## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 256 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 241
Hood | 02-92B/P | 200
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1014.6 | hPa |  | PRC PRO 01
Humidity | 56.6 | %/Hr
Internal temperature of calibration bench | 23.73 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1020 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 25 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.06 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.04 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3062 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 26/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 11/07/2018
Technical Manager | Anthony HUE | Approval |  | 16/07/2017
CEO | Robert GLEMOT | Approval |  | 16/07/2017
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3826 | 57.5422
0.012209 | -35.9329 | 38.4877
0.018313 | -34.8882 | 28.5029
0.024417 | -34.4392 | 22.1266
0.030521 | -34.3265 | 18.3276
0.036626 | -34.1847 | 15.2364
0.04273 | -34.1017 | 12.9381
0.048834 | -34.0486 | 11.3502
0.054938 | -33.979 | 10.1396
0.061043 | -33.9058 | 9.0168
0.067147 | -33.8947 | 8.1585
0.073251 | -33.8945 | 7.4117
0.079355 | -33.9103 | 6.8539
0.08546 | -33.9262 | 6.3937
0.091564 | -33.9219 | 6.013
0.097668 | -33.8884 | 5.7473
0.10377 | -33.8427 | 5.4009
0.10988 | -33.8131 | 5.0399
0.11598 | -33.8044 | 4.6832
0.12209 | -33.8205 | 4.4044
0.12819 | -33.8518 | 4.1706
0.13429 | -33.882 | 3.9313
0.1404 | -33.8981 | 3.6648
0.1465 | -33.9129 | 3.4119
0.15261 | -33.9196 | 3.1768
0.15871 | -33.9162 | 2.9819
0.16482 | -33.9022 | 2.8219
0.17092 | -33.8781 | 2.7545
0.17702 | -33.8595 | 2.6542
0.18313 | -33.8426 | 2.5692
0.18923 | -33.8305 | 2.4638
0.19534 | -33.8239 | 2.3624
0.20144 | -33.823 | 2.253
0.20754 | -33.8259 | 2.1814
0.21365 | -33.8282 | 2.1084
0.21975 | -33.8324 | 2.0443
0.22586 | -33.8401 | 1.9792
0.23196 | -33.8483 | 1.906
0.23807 | -33.8564 | 1.8276
0.24417 | -33.8581 | 1.7342
0.25027 | -33.8609 | 1.6499
0.25638 | -33.8606 | 1.5583
0.26248 | -33.8584 | 1.4722
0.26859 | -33.8559 | 1.3902
0.27469 | -33.8551 | 1.3055
0.2808 | -33.8608 | 1.2431
0.2869 | -33.8652 | 1.1763
0.293 | -33.8705 | 1.1202
0.29911 | -33.8766 | 1.0821
0.30521 | -33.8825 | 1.0456
0.31132 | -33.8864 | 1.0303
0.31742 | -33.8905 | 1.0071
0.32353 | -33.8938 | 0.98751
0.32963 | -33.8966 | 0.97073
0.33573 | -33.8983 | 0.95566
0.34184 | -33.9006 | 0.93748
0.34794 | -33.9035 | 0.9017
0.35405 | -33.9065 | 0.87076
0.36015 | -33.9098 | 0.83668
0.36626 | -33.9119 | 0.79935
0.37236 | -33.9141 | 0.75021
0.37846 | -33.916 | 0.69433
0.38457 | -33.9149 | 0.6373
0.39067 | -33.9154 | 0.58827
0.39678 | -33.9136 | 0.53359
0.40288 | -33.9112 | 0.48202
0.40899 | -33.9074 | 0.43155
0.41509 | -33.903 | 0.38079
0.42119 | -33.8979 | 0.33422
0.4273 | -33.8942 | 0.28628
0.4334 | -33.8907 | 0.2376
0.43951 | -33.8876 | 0.18961
0.44561 | -33.886 | 0.14453
0.45172 | -33.8873 | 0.10702
0.45782 | -33.8869 | 0.065558
0.46392 | -33.888 | 0.027517
0.47003 | -33.8895 | -0.007244
0.47613 | -33.8915 | -0.038835
0.48224 | -33.8936 | -0.068736
0.48834 | -33.896 | -0.089507
0.49445 | -33.8981 | -0.11365
0.50055 | -33.9006 | -0.13554
0.50665 | -33.9024 | -0.15316
0.51276 | -33.9036 | -0.17221
0.51886 | -33.9049 | -0.18878
0.52497 | -33.904 | -0.19514
0.53107 | -33.9045 | -0.20496
0.53717 | -33.9049 | -0.21574
0.54328 | -33.9048 | -0.2209
0.54938 | -33.9048 | -0.22896
0.55549 | -33.9043 | -0.2379
0.56159 | -33.9041 | -0.24952
0.5677 | -33.9037 | -0.26036
0.5738 | -33.9032 | -0.27383
0.5799 | -33.9032 | -0.28626
0.58601 | -33.9035 | -0.30337
0.59211 | -33.9045 | -0.32671
0.59822 | -33.905 | -0.34408
0.60432 | -33.9059 | -0.36611
0.61043 | -33.9068 | -0.38805
0.61653 | -33.9079 | -0.41014
0.62263 | -33.9089 | -0.43378
0.62874 | -33.9102 | -0.45462
0.63484 | -33.9112 | -0.47638
0.64095 | -33.9124 | -0.49797
0.64705 | -33.9135 | -0.51852
0.65316 | -33.9145 | -0.53843
0.65926 | -33.9155 | -0.56106
0.66536 | -33.915 | -0.57936
0.67147 | -33.9153 | -0.59601
0.67757 | -33.9158 | -0.61418
0.68368 | -33.9157 | -0.63402
0.68978 | -33.9158 | -0.65203
0.69589 | -33.9155 | -0.67016
0.70199 | -33.9148 | -0.69143
0.70809 | -33.9142 | -0.71137
0.7142 | -33.9138 | -0.73028
0.7203 | -33.913 | -0.74956
0.72641 | -33.9122 | -0.76989
0.73251 | -33.912 | -0.79026
0.73862 | -33.912 | -0.81398
0.74472 | -33.9122 | -0.83291
0.75082 | -33.9118 | -0.8534
0.75693 | -33.9119 | -0.87631
0.76303 | -33.9118 | -0.89827
0.76914 | -33.9121 | -0.92478
0.77524 | -33.9123 | -0.94927
0.78135 | -33.9126 | -0.97219
0.78745 | -33.9131 | -0.99661
0.79355 | -33.9134 | -1.0208
0.79966 | -33.9143 | -1.0462
0.80576 | -33.9153 | -1.076
0.81187 | -33.9155 | -1.1049
0.81797 | -33.9159 | -1.135
0.82408 | -33.9165 | -1.1617
0.83018 | -33.9169 | -1.1923
0.83628 | -33.9176 | -1.2223
0.84239 | -33.9174 | -1.2498
0.84849 | -33.9173 | -1.2773
0.8546 | -33.9176 | -1.3044
0.8607 | -33.9175 | -1.3331
0.86681 | -33.9176 | -1.3609
0.87291 | -33.9171 | -1.3862
0.87901 | -33.9165 | -1.4082
0.88512 | -33.9163 | -1.4308
0.89122 | -33.9161 | -1.4537
0.89733 | -33.9158 | -1.4728
0.90343 | -33.9155 | -1.4937
0.90953 | -33.9156 | -1.5085
0.91564 | -33.915 | -1.5243
0.92174 | -33.9147 | -1.5384
0.92785 | -33.9141 | -1.5518
0.93395 | -33.9141 | -1.5635
0.94006 | -33.9141 | -1.5772
0.94616 | -33.9139 | -1.582
0.95226 | -33.914 | -1.5935
0.95837 | -33.9141 | -1.6051
0.96447 | -33.9145 | -1.6154
0.97058 | -33.9145 | -1.6253
0.97668 | -33.9142 | -1.6333
0.98279 | -33.9143 | -1.637
0.98889 | -33.9145 | -1.6421
0.99499 | -33.9146 | -1.6506
1.0011 | -33.9148 | -1.6557
1.0072 | -33.9147 | -1.6596
1.0133 | -33.9147 | -1.6637
1.0194 | -33.9147 | -1.6691
1.0255 | -33.9144 | -1.676
1.0316 | -33.9138 | -1.6802
1.0377 | -33.9134 | -1.6861
1.0438 | -33.9126 | -1.6905
1.0499 | -33.9112 | -1.6968
1.056 | -33.9101 | -1.7016
1.0621 | -33.909 | -1.7073
1.0682 | -33.9085 | -1.7123
1.0743 | -33.9074 | -1.7199
1.0805 | -33.9068 | -1.7266
1.0866 | -33.9052 | -1.7372
1.0927 | -33.9039 | -1.7445
1.0988 | -33.9032 | -1.7549
1.1049 | -33.9018 | -1.7623
1.111 | -33.9002 | -1.7732
1.1171 | -33.8989 | -1.7824
1.1232 | -33.8976 | -1.7963
1.1293 | -33.8965 | -1.808
1.1354 | -33.8955 | -1.8202
1.1415 | -33.8943 | -1.8339
1.1476 | -33.893 | -1.8476
1.1537 | -33.8917 | -1.8606
1.1598 | -33.8915 | -1.8786
1.1659 | -33.891 | -1.8904
1.172 | -33.8903 | -1.9035
1.1781 | -33.8896 | -1.9173
1.1842 | -33.889 | -1.9326
1.1903 | -33.8897 | -1.9477
1.1964 | -33.8895 | -1.962
1.2025 | -33.8894 | -1.977
1.2086 | -33.8899 | -1.9926
1.2147 | -33.8901 | -2.0074
1.2209 | -33.8907 | -2.0208
1.227 | -33.8925 | -2.0358
1.2331 | -33.8935 | -2.0499
1.2392 | -33.8944 | -2.0656
1.2453 | -33.8955 | -2.0776
1.2514 | -33.8968 | -2.0898
1.2575 | -33.8983 | -2.1038
1.2636 | -33.9006 | -2.1153
1.2697 | -33.9022 | -2.1271
1.2758 | -33.9037 | -2.1411
1.2819 | -33.9054 | -2.1546
1.288 | -33.9071 | -2.1684
1.2941 | -33.909 | -2.1792
1.3002 | -33.911 | -2.1876
1.3063 | -33.9126 | -2.1977
1.3124 | -33.9143 | -2.2078
1.3185 | -33.9162 | -2.2186
1.3246 | -33.918 | -2.2294
1.3307 | -33.9201 | -2.2391
1.3368 | -33.9218 | -2.2467
1.3429 | -33.9234 | -2.2556
1.349 | -33.9248 | -2.2643
1.3551 | -33.9265 | -2.2753
1.3613 | -33.928 | -2.2833
1.3674 | -33.9292 | -2.2919
1.3735 | -33.9304 | -2.3
1.3796 | -33.9317 | -2.3076
1.3857 | -33.9326 | -2.3147
1.3918 | -33.9337 | -2.3245
1.3979 | -33.9349 | -2.333
1.404 | -33.9355 | -2.343
1.4101 | -33.9364 | -2.3521
1.4162 | -33.9372 | -2.3607
1.4223 | -33.938 | -2.3709
1.4284 | -33.9387 | -2.3793
1.4345 | -33.9393 | -2.3896
1.4406 | -33.9397 | -2.4006
1.4467 | -33.9404 | -2.41
1.4528 | -33.941 | -2.4192
1.4589 | -33.9418 | -2.4297
1.465 | -33.942 | -2.4374
1.4711 | -33.9429 | -2.4464
1.4772 | -33.9433 | -2.4584
1.4833 | -33.9439 | -2.4705
1.4894 | -33.9446 | -2.4818
1.4955 | -33.9453 | -2.4933
1.5016 | -33.9459 | -2.5052
1.5078 | -33.9463 | -2.5195
1.5139 | -33.9469 | -2.532
1.52 | -33.9474 | -2.5451
1.5261 | -33.9478 | -2.5566
1.5322 | -33.948 | -2.5691
1.5383 | -33.9482 | -2.5811
1.5444 | -33.9488 | -2.5943
1.5505 | -33.9492 | -2.6073
1.5566 | -33.9497 | -2.62
1.5627 | -33.9502 | -2.6331
1.5688 | -33.951 | -2.6461
1.5749 | -33.9512 | -2.6583
1.581 | -33.9515 | -2.6748
1.5871 | -33.952 | -2.6901
1.5932 | -33.9524 | -2.7028
1.5993 | -33.9527 | -2.7193
1.6054 | -33.9531 | -2.7341
1.6115 | -33.9534 | -2.7508
1.6176 | -33.9537 | -2.7669
1.6237 | -33.9539 | -2.7812
1.6298 | -33.9542 | -2.7964
1.6359 | -33.9546 | -2.8107
1.642 | -33.955 | -2.8251
1.6482 | -33.9554 | -2.8419
1.6543 | -33.9558 | -2.8584
1.6604 | -33.9562 | -2.8729
1.6665 | -33.9563 | -2.8876
1.6726 | -33.9566 | -2.9037
1.6787 | -33.9571 | -2.9205
1.6848 | -33.9569 | -2.9368
1.6909 | -33.9572 | -2.9528
1.697 | -33.9575 | -2.9688
1.7031 | -33.9575 | -2.9844
1.7092 | -33.9575 | -2.9991
1.7153 | -33.9577 | -3.0152
1.7214 | -33.9573 | -3.0305
1.7275 | -33.9571 | -3.0462
1.7336 | -33.9569 | -3.0622
1.7397 | -33.9569 | -3.0773
1.7458 | -33.957 | -3.0938
1.7519 | -33.9568 | -3.1098
1.758 | -33.9567 | -3.1246
1.7641 | -33.9568 | -3.139
1.7702 | -33.9565 | -3.1546
1.7763 | -33.9566 | -3.1682
1.7824 | -33.9564 | -3.1842
1.7885 | -33.9562 | -3.1982
1.7947 | -33.9563 | -3.2129
1.8008 | -33.9562 | -3.2258
1.8069 | -33.9561 | -3.2393
1.813 | -33.9561 | -3.2534
1.8191 | -33.9558 | -3.2675
1.8252 | -33.9556 | -3.2813
1.8313 | -33.9556 | -3.2956
1.8374 | -33.9556 | -3.309
1.8435 | -33.9556 | -3.3211
1.8496 | -33.9553 | -3.3354
1.8557 | -33.9552 | -3.3501
1.8618 | -33.9552 | -3.3618
1.8679 | -33.9552 | -3.3747
1.874 | -33.9552 | -3.3878
1.8801 | -33.955 | -3.4006
1.8862 | -33.9551 | -3.4133
1.8923 | -33.955 | -3.4254
1.8984 | -33.9547 | -3.436
1.9045 | -33.9546 | -3.447
1.9106 | -33.9543 | -3.4588
1.9167 | -33.954 | -3.47
1.9228 | -33.954 | -3.4808
1.9289 | -33.9537 | -3.4896
1.9351 | -33.9537 | -3.4987
1.9412 | -33.9537 | -3.509
1.9473 | -33.9538 | -3.5209
1.9534 | -33.9537 | -3.5319
1.9595 | -33.9534 | -3.5433
1.9656 | -33.953 | -3.5539
1.9717 | -33.9527 | -3.5644
1.9778 | -33.9527 | -3.5759
1.9839 | -33.9527 | -3.5872
1.99 | -33.9527 | -3.5971
1.9961 | -33.9524 | -3.6074
2.0022 | -33.9523 | -3.6174
2.0083 | -33.952 | -3.6273
2.0144 | -33.952 | -3.6371
2.0205 | -33.952 | -3.6488
2.0266 | -33.952 | -3.6603
2.0327 | -33.9521 | -3.6714
2.0388 | -33.9522 | -3.6852
2.0449 | -33.9523 | -3.6954
2.051 | -33.9523 | -3.7054
2.0571 | -33.9522 | -3.7167
2.0632 | -33.952 | -3.7272
2.0693 | -33.952 | -3.7382
2.0754 | -33.9522 | -3.7499
2.0816 | -33.9523 | -3.7618
2.0877 | -33.9522 | -3.7728
2.0938 | -33.9524 | -3.7851
2.0999 | -33.9523 | -3.7966
2.106 | -33.9526 | -3.8084
2.1121 | -33.9525 | -3.8206
2.1182 | -33.9526 | -3.8306
2.1243 | -33.9528 | -3.8411
2.1304 | -33.9529 | -3.8514
2.1365 | -33.9532 | -3.8628
2.1426 | -33.9533 | -3.876
2.1487 | -33.9537 | -3.8876
2.1548 | -33.9541 | -3.9005
2.1609 | -33.9546 | -3.9124
2.167 | -33.955 | -3.9249
2.1731 | -33.9553 | -3.9367
2.1792 | -33.956 | -3.9479
2.1853 | -33.9565 | -3.9606
2.1914 | -33.9569 | -3.9727
2.1975 | -33.9574 | -3.9847
2.2036 | -33.9578 | -3.9978
2.2097 | -33.9582 | -4.0097
2.2158 | -33.9591 | -4.02
2.222 | -33.9597 | -4.0301
2.2281 | -33.96 | -4.0408
2.2342 | -33.9604 | -4.0538
2.2403 | -33.9609 | -4.0639
2.2464 | -33.9616 | -4.0749
2.2525 | -33.962 | -4.0856
2.2586 | -33.9629 | -4.0955
2.2647 | -33.9637 | -4.1048
2.2708 | -33.9642 | -4.1164
2.2769 | -33.9651 | -4.1279
2.283 | -33.9655 | -4.1372
2.2891 | -33.966 | -4.1476
2.2952 | -33.9666 | -4.1573
2.3013 | -33.9672 | -4.1675
2.3074 | -33.9679 | -4.1771
2.3135 | -33.9683 | -4.1884
2.3196 | -33.9689 | -4.197
2.3257 | -33.9693 | -4.2076
2.3318 | -33.9696 | -4.2166
2.3379 | -33.9701 | -4.228
2.344 | -33.9704 | -4.2388
2.3501 | -33.9708 | -4.2495
2.3562 | -33.9711 | -4.2579
2.3623 | -33.9714 | -4.2666
2.3685 | -33.9716 | -4.2757
2.3746 | -33.9719 | -4.2862
2.3807 | -33.9724 | -4.2951
2.3868 | -33.9725 | -4.3042
2.3929 | -33.9729 | -4.3135
2.399 | -33.9732 | -4.3216
2.4051 | -33.9737 | -4.3311
2.4112 | -33.974 | -4.3401
2.4173 | -33.9743 | -4.3491
2.4234 | -33.9743 | -4.3577
2.4295 | -33.9745 | -4.3663
2.4356 | -33.9746 | -4.3756
2.4417 | -33.9747 | -4.3857
2.4478 | -33.9748 | -4.3952
2.4539 | -33.9751 | -4.4032
2.46 | -33.9747 | -4.4129
2.4661 | -33.9748 | -4.4222
2.4722 | -33.975 | -4.4309
2.4783 | -33.9751 | -4.4394
2.4844 | -33.9751 | -4.448
2.4905 | -33.975 | -4.4572
2.4966 | -33.9748 | -4.4662
2.5027 | -33.9747 | -4.4743
2.5089 | -33.9746 | -4.4828
2.515 | -33.9744 | -4.4921
2.5211 | -33.9744 | -4.5005
2.5272 | -33.974 | -4.5096
2.5333 | -33.9739 | -4.5192
2.5394 | -33.9738 | -4.5282
2.5455 | -33.9736 | -4.5377
2.5516 | -33.9735 | -4.5465
2.5577 | -33.9735 | -4.5564
2.5638 | -33.9732 | -4.5667
2.5699 | -33.973 | -4.577
2.576 | -33.9728 | -4.5874
2.5821 | -33.9726 | -4.5974
2.5882 | -33.9724 | -4.607
2.5943 | -33.9722 | -4.6168
2.6004 | -33.9721 | -4.6268
2.6065 | -33.9719 | -4.6353
2.6126 | -33.9718 | -4.6445
2.6187 | -33.9717 | -4.6533
2.6248 | -33.9715 | -4.6624
2.6309 | -33.9713 | -4.6737
2.637 | -33.9714 | -4.6849
2.6431 | -33.9713 | -4.6945
2.6492 | -33.9712 | -4.7045
2.6554 | -33.9712 | -4.7148
2.6615 | -33.9709 | -4.7247
2.6676 | -33.9709 | -4.7352
2.6737 | -33.971 | -4.7459
2.6798 | -33.971 | -4.7559
2.6859 | -33.971 | -4.7664
2.692 | -33.9709 | -4.7762
2.6981 | -33.9708 | -4.7874
2.7042 | -33.9709 | -4.7994
2.7103 | -33.9708 | -4.8096
2.7164 | -33.971 | -4.8203
2.7225 | -33.9711 | -4.8316
2.7286 | -33.9712 | -4.8428
2.7347 | -33.9712 | -4.8529
2.7408 | -33.9715 | -4.8621
2.7469 | -33.9714 | -4.8736
2.753 | -33.9717 | -4.8842
2.7591 | -33.9718 | -4.8962
2.7652 | -33.9718 | -4.9077
2.7713 | -33.9718 | -4.9176
2.7774 | -33.9722 | -4.929
2.7835 | -33.9723 | -4.9397
2.7896 | -33.9724 | -4.9491
2.7958 | -33.9728 | -4.9595
2.8019 | -33.9733 | -4.9693
2.808 | -33.9735 | -4.9812
2.8141 | -33.9739 | -4.9941
2.8202 | -33.9743 | -5.0063
2.8263 | -33.9747 | -5.0172
2.8324 | -33.9749 | -5.0296
2.8385 | -33.9753 | -5.0417
2.8446 | -33.9758 | -5.0534
2.8507 | -33.9762 | -5.0655
2.8568 | -33.9766 | -5.0767
2.8629 | -33.977 | -5.088
2.869 | -33.9774 | -5.0993
2.8751 | -33.9778 | -5.1113
2.8812 | -33.9784 | -5.1245
2.8873 | -33.9788 | -5.1358
2.8934 | -33.9791 | -5.1477
2.8995 | -33.9794 | -5.1597
2.9056 | -33.9797 | -5.1704
2.9117 | -33.9802 | -5.182
2.9178 | -33.9808 | -5.1946
2.9239 | -33.9812 | -5.2079
2.93 | -33.9817 | -5.2213
2.9361 | -33.9822 | -5.2327
2.9423 | -33.9828 | -5.2447
2.9484 | -33.9832 | -5.2582
2.9545 | -33.9839 | -5.2714
2.9606 | -33.9844 | -5.2847
2.9667 | -33.9849 | -5.2965
2.9728 | -33.9852 | -5.3087
2.9789 | -33.9857 | -5.3214
2.985 | -33.9861 | -5.335
2.9911 | -33.9865 | -5.3475
2.9972 | -33.9871 | -5.3602
3.0033 | -33.9876 | -5.3732
3.0094 | -33.9881 | -5.3862
3.0155 | -33.9887 | -5.3987
3.0216 | -33.9892 | -5.4121
3.0277 | -33.9896 | -5.4257
3.0338 | -33.9902 | -5.4392
3.0399 | -33.9907 | -5.4528
3.046 | -33.991 | -5.4658
3.0521 | -33.9916 | -5.4782
3.0582 | -33.9918 | -5.4919
3.0643 | -33.9924 | -5.5049
3.0704 | -33.9926 | -5.5182
3.0765 | -33.993 | -5.5304
3.0827 | -33.9932 | -5.5434
3.0888 | -33.9936 | -5.5562
3.0949 | -33.9939 | -5.5691
3.101 | -33.9943 | -5.5831
3.1071 | -33.9946 | -5.5971
3.1132 | -33.995 | -5.611
3.1193 | -33.9954 | -5.624
3.1254 | -33.9958 | -5.6361
3.1315 | -33.9962 | -5.6503
3.1376 | -33.9964 | -5.6631
3.1437 | -33.9967 | -5.6766
3.1498 | -33.9971 | -5.6906
3.1559 | -33.9974 | -5.7037
3.162 | -33.9974 | -5.7177
3.1681 | -33.9975 | -5.7317
3.1742 | -33.9978 | -5.7454
3.1803 | -33.9982 | -5.7584
3.1864 | -33.9985 | -5.7706
3.1925 | -33.9988 | -5.7842
3.1986 | -33.9989 | -5.7969
3.2047 | -33.999 | -5.8108
3.2108 | -33.9992 | -5.8224
3.2169 | -33.9996 | -5.8361
3.223 | -33.9995 | -5.8493
3.2292 | -33.9995 | -5.8632
3.2353 | -33.9994 | -5.8727
3.2414 | -33.9995 | -5.8856
3.2475 | -33.9995 | -5.8965
3.2536 | -33.9998 | -5.9102
3.2597 | -34 | -5.9213
3.2658 | -34.0001 | -5.932
3.2719 | -34.0003 | -5.9431
3.278 | -34.0003 | -5.9551
3.2841 | -34.0004 | -5.967
3.2902 | -34.0005 | -5.9794
3.2963 | -34.0007 | -5.991
3.3024 | -34.0005 | -6.0012
3.3085 | -34.0008 | -6.0136
3.3146 | -34.0006 | -6.0245
3.3207 | -34.0008 | -6.0366
3.3268 | -34.0009 | -6.0487
3.3329 | -34.001 | -6.0585
3.339 | -34.0011 | -6.0682
3.3451 | -34.0011 | -6.0795
3.3512 | -34.0012 | -6.0889
3.3573 | -34.0014 | -6.1002
3.3634 | -34.0015 | -6.1117
3.3696 | -34.0014 | -6.1235
3.3757 | -34.0014 | -6.1333
3.3818 | -34.0017 | -6.1449
3.3879 | -34.0016 | -6.1553
3.394 | -34.0016 | -6.1657
3.4001 | -34.0016 | -6.1766
3.4062 | -34.0016 | -6.1867
3.4123 | -34.0016 | -6.1965
3.4184 | -34.0017 | -6.2074
3.4245 | -34.0018 | -6.2179
3.4306 | -34.0017 | -6.2277
3.4367 | -34.0018 | -6.2374
3.4428 | -34.0021 | -6.2466
3.4489 | -34.002 | -6.2577
3.455 | -34.0019 | -6.2681
3.4611 | -34.0019 | -6.2791
3.4672 | -34.0019 | -6.2904
3.4733 | -34.002 | -6.3005
3.4794 | -34.002 | -6.311
3.4855 | -34.0022 | -6.3206
3.4916 | -34.0023 | -6.3303
3.4977 | -34.0023 | -6.3412
3.5038 | -34.0024 | -6.3511
3.5099 | -34.0025 | -6.3607
3.5161 | -34.0026 | -6.3702
3.5222 | -34.0027 | -6.3802
3.5283 | -34.0026 | -6.3903
3.5344 | -34.0026 | -6.401
3.5405 | -34.0027 | -6.4109
3.5466 | -34.0029 | -6.4206
3.5527 | -34.0029 | -6.4313
3.5588 | -34.003 | -6.441
3.5649 | -34.0031 | -6.4512
3.571 | -34.0031 | -6.4614
3.5771 | -34.0032 | -6.4721
3.5832 | -34.0034 | -6.4817
3.5893 | -34.0036 | -6.4915
3.5954 | -34.0037 | -6.502
3.6015 | -34.0038 | -6.5123
3.6076 | -34.004 | -6.5219
3.6137 | -34.0041 | -6.5318
3.6198 | -34.0044 | -6.5408
3.6259 | -34.0044 | -6.5515
3.632 | -34.0047 | -6.5621
3.6381 | -34.0049 | -6.5719
3.6442 | -34.0051 | -6.5823
3.6503 | -34.0053 | -6.5914
3.6565 | -34.0056 | -6.6031
3.6626 | -34.0056 | -6.6138
3.6687 | -34.0059 | -6.6243
3.6748 | -34.0062 | -6.6342
3.6809 | -34.0066 | -6.6447
3.687 | -34.0067 | -6.6558
3.6931 | -34.0071 | -6.6658
3.6992 | -34.0073 | -6.6764
3.7053 | -34.0075 | -6.6864
3.7114 | -34.0077 | -6.6976
3.7175 | -34.008 | -6.7075
3.7236 | -34.0082 | -6.7163
3.7297 | -34.0083 | -6.727
3.7358 | -34.0086 | -6.7386
3.7419 | -34.0087 | -6.7501
3.748 | -34.0089 | -6.7608
3.7541 | -34.009 | -6.7717
3.7602 | -34.0094 | -6.7825
3.7663 | -34.0096 | -6.7927
3.7724 | -34.0099 | -6.8036
3.7785 | -34.0102 | -6.814
3.7846 | -34.0104 | -6.8246
3.7907 | -34.0108 | -6.8355
3.7969 | -34.011 | -6.847
3.803 | -34.0113 | -6.857
3.8091 | -34.0116 | -6.8673
3.8152 | -34.0119 | -6.8785
3.8213 | -34.012 | -6.89
3.8274 | -34.0124 | -6.9005
3.8335 | -34.0128 | -6.912
3.8396 | -34.0132 | -6.9234
3.8457 | -34.0136 | -6.9336
3.8518 | -34.014 | -6.9445
3.8579 | -34.0144 | -6.955
3.864 | -34.0147 | -6.9657
3.8701 | -34.0151 | -6.9766
3.8762 | -34.0154 | -6.9877
3.8823 | -34.0157 | -6.9983
3.8884 | -34.0161 | -7.0086
3.8945 | -34.0165 | -7.0199
3.9006 | -34.0169 | -7.0314
3.9067 | -34.0172 | -7.0432
3.9128 | -34.0176 | -7.0544
3.9189 | -34.018 | -7.065
3.925 | -34.0182 | -7.0758
3.9311 | -34.0186 | -7.0858
3.9372 | -34.019 | -7.0973
3.9434 | -34.0193 | -7.1077
3.9495 | -34.0197 | -7.1193
3.9556 | -34.0201 | -7.1304
3.9617 | -34.0204 | -7.1413
3.9678 | -34.0208 | -7.152
3.9739 | -34.0212 | -7.1632
3.98 | -34.0215 | -7.1749
3.9861 | -34.0218 | -7.1854
3.9922 | -34.022 | -7.197
3.9983 | -34.0224 | -7.207
4.0044 | -34.0227 | -7.2167
4.0105 | -34.023 | -7.2276
4.0166 | -34.0235 | -7.2386
4.0227 | -34.0239 | -7.2489
4.0288 | -34.0241 | -7.2603
4.0349 | -34.0244 | -7.2701
4.041 | -34.0247 | -7.281
4.0471 | -34.025 | -7.2913
4.0532 | -34.0254 | -7.3024
4.0593 | -34.0257 | -7.3139
4.0654 | -34.026 | -7.3247
4.0715 | -34.0263 | -7.336
4.0776 | -34.0266 | -7.3464
4.0838 | -34.027 | -7.3577
4.0899 | -34.0273 | -7.3689
4.096 | -34.0276 | -7.3793
4.1021 | -34.0279 | -7.3901
4.1082 | -34.0283 | -7.4004
4.1143 | -34.0286 | -7.4104
4.1204 | -34.029 | -7.4211
4.1265 | -34.0294 | -7.4315
4.1326 | -34.0297 | -7.4418
4.1387 | -34.0301 | -7.453
4.1448 | -34.0304 | -7.4639
4.1509 | -34.0306 | -7.4746
4.157 | -34.0309 | -7.4858
4.1631 | -34.0312 | -7.497
4.1692 | -34.0315 | -7.5076
4.1753 | -34.0319 | -7.5185
4.1814 | -34.0321 | -7.5283
4.1875 | -34.0324 | -7.5397
4.1936 | -34.0327 | -7.551
4.1997 | -34.033 | -7.5611
4.2058 | -34.0332 | -7.5726
4.2119 | -34.0335 | -7.5823
4.218 | -34.0338 | -7.5925
4.2241 | -34.0341 | -7.6021
4.2303 | -34.0345 | -7.6133
4.2364 | -34.0348 | -7.6247
4.2425 | -34.0351 | -7.6352
4.2486 | -34.0353 | -7.6463
4.2547 | -34.0354 | -7.6556
4.2608 | -34.0357 | -7.6655
4.2669 | -34.036 | -7.6759
4.273 | -34.0363 | -7.6867
4.2791 | -34.0365 | -7.6974
4.2852 | -34.0366 | -7.7077
4.2913 | -34.0368 | -7.7169
4.2974 | -34.037 | -7.7279
4.3035 | -34.0372 | -7.7386
4.3096 | -34.0373 | -7.7491
4.3157 | -34.0376 | -7.759
4.3218 | -34.0378 | -7.769
4.3279 | -34.038 | -7.7794
4.334 | -34.0382 | -7.7892
4.3401 | -34.0384 | -7.7992
4.3462 | -34.0387 | -7.81
4.3523 | -34.039 | -7.8197
4.3584 | -34.0392 | -7.8282
4.3645 | -34.0393 | -7.8379
4.3707 | -34.0395 | -7.8472
4.3768 | -34.0397 | -7.8576
4.3829 | -34.0399 | -7.8679
4.389 | -34.0401 | -7.8776
4.3951 | -34.0404 | -7.8872
4.4012 | -34.0405 | -7.897
4.4073 | -34.0408 | -7.9072
4.4134 | -34.041 | -7.9173
4.4195 | -34.0412 | -7.927
4.4256 | -34.0415 | -7.9366
4.4317 | -34.0417 | -7.9461
4.4378 | -34.0418 | -7.9557
4.4439 | -34.042 | -7.9652
4.45 | -34.0421 | -7.9752
4.4561 | -34.0422 | -7.985
4.4622 | -34.0423 | -7.9947
4.4683 | -34.0425 | -8.0039
4.4744 | -34.0427 | -8.0141
4.4805 | -34.0429 | -8.0232
4.4866 | -34.0431 | -8.0325
4.4927 | -34.0433 | -8.0427
4.4988 | -34.0435 | -8.0521
4.5049 | -34.0437 | -8.0613
4.511 | -34.0439 | -8.0704
4.5172 | -34.0442 | -8.0802
4.5233 | -34.0444 | -8.0908
4.5294 | -34.0446 | -8.1002
4.5355 | -34.0448 | -8.1096
4.5416 | -34.0451 | -8.1187
4.5477 | -34.0452 | -8.128
4.5538 | -34.0454 | -8.1368
4.5599 | -34.0457 | -8.1463
4.566 | -34.0459 | -8.1561
4.5721 | -34.0461 | -8.1653
4.5782 | -34.0464 | -8.1751
4.5843 | -34.0466 | -8.1847
4.5904 | -34.0469 | -8.1941
4.5965 | -34.0472 | -8.2028
4.6026 | -34.0474 | -8.2134
4.6087 | -34.0476 | -8.2225
4.6148 | -34.0478 | -8.2331
4.6209 | -34.0482 | -8.2416
4.627 | -34.0485 | -8.2517
4.6331 | -34.0485 | -8.2606
4.6392 | -34.0488 | -8.2699
4.6453 | -34.049 | -8.2793
4.6514 | -34.0493 | -8.2888
4.6576 | -34.0493 | -8.2993
4.6637 | -34.0495 | -8.307
4.6698 | -34.0497 | -8.3156
4.6759 | -34.0498 | -8.3254
4.682 | -34.05 | -8.3345
4.6881 | -34.0502 | -8.3434
4.6942 | -34.0504 | -8.3527
4.7003 | -34.0506 | -8.3621
4.7064 | -34.0509 | -8.3714
4.7125 | -34.0512 | -8.3809
4.7186 | -34.0513 | -8.3894
4.7247 | -34.0516 | -8.399
4.7308 | -34.0519 | -8.4083
4.7369 | -34.0521 | -8.4179
4.743 | -34.0523 | -8.4282
4.7491 | -34.0525 | -8.4379
4.7552 | -34.0526 | -8.4471
4.7613 | -34.0528 | -8.4568
4.7674 | -34.053 | -8.4654
4.7735 | -34.0531 | -8.4755
4.7796 | -34.0534 | -8.4852
4.7857 | -34.0535 | -8.4949
4.7918 | -34.0537 | -8.5036
4.7979 | -34.054 | -8.5127
4.8041 | -34.0541 | -8.522
4.8102 | -34.0544 | -8.5311
4.8163 | -34.0547 | -8.5406
4.8224 | -34.0548 | -8.5493
4.8285 | -34.0551 | -8.5584
4.8346 | -34.0553 | -8.568
4.8407 | -34.0555 | -8.5773
4.8468 | -34.0557 | -8.587
4.8529 | -34.0559 | -8.5968
4.859 | -34.0561 | -8.6063
4.8651 | -34.0564 | -8.6157
4.8712 | -34.0566 | -8.6251
4.8773 | -34.0568 | -8.6347
4.8834 | -34.057 | -8.6444
4.8895 | -34.0573 | -8.6545
4.8956 | -34.0575 | -8.6644
4.9017 | -34.0576 | -8.6739
4.9078 | -34.0578 | -8.6838
4.9139 | -34.0581 | -8.6924
4.92 | -34.0583 | -8.7014
4.9261 | -34.0586 | -8.7103
4.9322 | -34.059 | -8.7195
4.9383 | -34.0593 | -8.7286
4.9445 | -34.0595 | -8.7387
4.9506 | -34.0597 | -8.7476
4.9567 | -34.06 | -8.7569
4.9628 | -34.0603 | -8.7663
4.9689 | -34.0604 | -8.7761
4.975 | -34.0606 | -8.7855
4.9811 | -34.0609 | -8.7953
4.9872 | -34.0611 | -8.8051
4.9933 | -34.0615 | -8.8148
4.9994 | -34.0618 | -8.8246
5.0055 | -34.0621 | -8.8343
5.0116 | -34.0625 | -8.8445
5.0177 | -34.0628 | -8.8546
5.0238 | -34.0631 | -8.865
5.0299 | -34.0636 | -8.876
5.036 | -34.0638 | -8.8856
5.0421 | -34.064 | -8.8953
5.0482 | -34.0643 | -8.9056
5.0543 | -34.0646 | -8.9149
5.0604 | -34.065 | -8.9252
5.0665 | -34.0653 | -8.9349
5.0726 | -34.0656 | -8.9453
5.0787 | -34.0659 | -8.9555
5.0848 | -34.0662 | -8.9655
5.091 | -34.0665 | -8.9754
5.0971 | -34.0668 | -8.9855
5.1032 | -34.0671 | -8.9959
5.1093 | -34.0675 | -9.0058
5.1154 | -34.0677 | -9.0158
5.1215 | -34.0679 | -9.0259
5.1276 | -34.0682 | -9.0357
5.1337 | -34.0687 | -9.0454
5.1398 | -34.0689 | -9.0552
5.1459 | -34.0693 | -9.0657
5.152 | -34.0696 | -9.0767
5.1581 | -34.0698 | -9.0873
5.1642 | -34.0701 | -9.0973
5.1703 | -34.0704 | -9.1073
5.1764 | -34.0707 | -9.1173
5.1825 | -34.071 | -9.1276
5.1886 | -34.0713 | -9.1374
5.1947 | -34.0716 | -9.1482
5.2008 | -34.072 | -9.1581
5.2069 | -34.0723 | -9.1692
5.213 | -34.0726 | -9.1788
5.2191 | -34.073 | -9.1891
5.2252 | -34.0733 | -9.1994
5.2314 | -34.0735 | -9.2102
5.2375 | -34.0739 | -9.2201
5.2436 | -34.0742 | -9.2303
5.2497 | -34.0745 | -9.2402
5.2558 | -34.0749 | -9.2511
5.2619 | -34.0752 | -9.262
5.268 | -34.0755 | -9.2724
5.2741 | -34.0759 | -9.2836
5.2802 | -34.0762 | -9.2944
5.2863 | -34.0766 | -9.3046
5.2924 | -34.0768 | -9.3157
5.2985 | -34.0773 | -9.327
5.3046 | -34.0777 | -9.3382
5.3107 | -34.0781 | -9.349
5.3168 | -34.0785 | -9.3595
5.3229 | -34.0788 | -9.3693
5.329 | -34.0792 | -9.3793
5.3351 | -34.0796 | -9.3895
5.3412 | -34.08 | -9.3996
5.3473 | -34.0804 | -9.4105
5.3534 | -34.0808 | -9.4213
5.3595 | -34.081 | -9.432
5.3656 | -34.0814 | -9.443
5.3717 | -34.0817 | -9.4535
5.3779 | -34.082 | -9.4644
5.384 | -34.0824 | -9.4751
5.3901 | -34.0827 | -9.4858
5.3962 | -34.083 | -9.4966
5.4023 | -34.0834 | -9.5073
5.4084 | -34.0838 | -9.5177
5.4145 | -34.0841 | -9.5286
5.4206 | -34.0843 | -9.5399
5.4267 | -34.0847 | -9.5512
5.4328 | -34.0851 | -9.5613
5.4389 | -34.0855 | -9.5716
5.445 | -34.0857 | -9.5829
5.4511 | -34.0861 | -9.5932
5.4572 | -34.0865 | -9.6044
5.4633 | -34.0868 | -9.6148
5.4694 | -34.0871 | -9.6252
5.4755 | -34.0875 | -9.6361
5.4816 | -34.0878 | -9.6467
5.4877 | -34.0881 | -9.6573
5.4938 | -34.0885 | -9.6677
5.4999 | -34.0888 | -9.6784
5.506 | -34.0892 | -9.6891
5.5121 | -34.0895 | -9.7002
5.5183 | -34.0898 | -9.7103
5.5244 | -34.0901 | -9.7211
5.5305 | -34.0904 | -9.7319
5.5366 | -34.0908 | -9.7424
5.5427 | -34.0911 | -9.7536
5.5488 | -34.0914 | -9.7646
5.5549 | -34.0917 | -9.7752
5.561 | -34.092 | -9.7865
5.5671 | -34.0924 | -9.7968
5.5732 | -34.0927 | -9.8077
5.5793 | -34.093 | -9.8192
5.5854 | -34.0935 | -9.8301
5.5915 | -34.0936 | -9.841
5.5976 | -34.094 | -9.852
5.6037 | -34.0943 | -9.8627
5.6098 | -34.0946 | -9.8732
5.6159 | -34.095 | -9.8841
5.622 | -34.0953 | -9.8945
5.6281 | -34.0956 | -9.9053
5.6342 | -34.096 | -9.9162
5.6403 | -34.0962 | -9.9264
5.6464 | -34.0967 | -9.9371
5.6525 | -34.0969 | -9.9476
5.6586 | -34.0973 | -9.9576
5.6648 | -34.0976 | -9.9688
5.6709 | -34.0979 | -9.9799
5.677 | -34.0983 | -9.9913
5.6831 | -34.0986 | -10.0016
5.6892 | -34.0989 | -10.0123
5.6953 | -34.0992 | -10.0229
5.7014 | -34.0995 | -10.0342
5.7075 | -34.0998 | -10.0451
5.7136 | -34.1002 | -10.0556
5.7197 | -34.1005 | -10.0662
5.7258 | -34.1007 | -10.0767
5.7319 | -34.101 | -10.0874
5.738 | -34.1014 | -10.0981
5.7441 | -34.1017 | -10.1088
5.7502 | -34.1021 | -10.1191
5.7563 | -34.1024 | -10.1295
5.7624 | -34.1027 | -10.1409
5.7685 | -34.103 | -10.1514
5.7746 | -34.1032 | -10.1619
5.7807 | -34.1035 | -10.1723
5.7868 | -34.1038 | -10.1829
5.7929 | -34.1041 | -10.1934
5.799 | -34.1044 | -10.204
5.8052 | -34.1047 | -10.2147
5.8113 | -34.105 | -10.2254
5.8174 | -34.1052 | -10.2362
5.8235 | -34.1055 | -10.2466
5.8296 | -34.1058 | -10.2578
5.8357 | -34.106 | -10.2685
5.8418 | -34.1064 | -10.2793
5.8479 | -34.1066 | -10.2902
5.854 | -34.1069 | -10.3011
5.8601 | -34.1072 | -10.3116
5.8662 | -34.1075 | -10.3225
5.8723 | -34.1077 | -10.3326
5.8784 | -34.108 | -10.3432
5.8845 | -34.1083 | -10.3539
5.8906 | -34.1086 | -10.3644
5.8967 | -34.1089 | -10.3747
5.9028 | -34.1093 | -10.3853
5.9089 | -34.1095 | -10.3953
5.915 | -34.1098 | -10.4059
5.9211 | -34.1101 | -10.4165
5.9272 | -34.1104 | -10.427
5.9333 | -34.1107 | -10.437
5.9394 | -34.111 | -10.4473
5.9455 | -34.1113 | -10.4575
5.9517 | -34.1117 | -10.468
5.9578 | -34.112 | -10.4789
5.9639 | -34.1124 | -10.4891
5.97 | -34.1127 | -10.4997
5.9761 | -34.113 | -10.51
5.9822 | -34.1133 | -10.5206
5.9883 | -34.1136 | -10.5309
5.9944 | -34.1139 | -10.5414
6.0005 | -34.1143 | -10.5522
6.0066 | -34.1146 | -10.5623
6.0127 | -34.1149 | -10.5721
6.0188 | -34.1152 | -10.583
6.0249 | -34.1155 | -10.5934
6.031 | -34.1158 | -10.6037
6.0371 | -34.1162 | -10.6144
6.0432 | -34.1165 | -10.6248
6.0493 | -34.1169 | -10.6351
6.0554 | -34.1172 | -10.6455
6.0615 | -34.1175 | -10.6564
6.0676 | -34.1178 | -10.6675
6.0737 | -34.1182 | -10.6781
6.0798 | -34.1184 | -10.6884
6.0859 | -34.1188 | -10.6988
6.0921 | -34.1191 | -10.709
6.0982 | -34.1195 | -10.7197
6.1043 | -34.1199 | -10.73
6.1104 | -34.1202 | -10.7409
6.1165 | -34.1205 | -10.7506
6.1226 | -34.1209 | -10.7605
6.1287 | -34.1212 | -10.771
6.1348 | -34.1216 | -10.7815
6.1409 | -34.1219 | -10.7918
6.147 | -34.1223 | -10.8021
6.1531 | -34.1226 | -10.8125
6.1592 | -34.1229 | -10.8229
6.1653 | -34.1232 | -10.833
6.1714 | -34.1235 | -10.8435
6.1775 | -34.1239 | -10.8541
6.1836 | -34.1242 | -10.8643
6.1897 | -34.1245 | -10.8747
6.1958 | -34.1248 | -10.8853
6.2019 | -34.1252 | -10.896
6.208 | -34.1255 | -10.9067
6.2141 | -34.1259 | -10.9173
6.2202 | -34.1262 | -10.9282
6.2263 | -34.1266 | -10.9384
6.2325 | -34.127 | -10.9483
6.2386 | -34.1274 | -10.9584
6.2447 | -34.1276 | -10.969
6.2508 | -34.128 | -10.9793
6.2569 | -34.1283 | -10.99
6.263 | -34.1287 | -11
6.2691 | -34.1291 | -11.0103
6.2752 | -34.1294 | -11.0205
6.2813 | -34.1297 | -11.0309
6.2874 | -34.1301 | -11.0413
6.2935 | -34.1305 | -11.0514
6.2996 | -34.1308 | -11.0619
6.3057 | -34.1312 | -11.0721
6.3118 | -34.1315 | -11.0824
6.3179 | -34.1318 | -11.0928
6.324 | -34.1322 | -11.1034
6.3301 | -34.1326 | -11.1135
6.3362 | -34.133 | -11.1239
6.3423 | -34.1334 | -11.1343
6.3484 | -34.1338 | -11.1448
6.3545 | -34.1341 | -11.1546
6.3606 | -34.1345 | -11.1646
6.3667 | -34.1349 | -11.1747
6.3728 | -34.1353 | -11.1851
6.379 | -34.1356 | -11.1954
6.3851 | -34.136 | -11.2058
6.3912 | -34.1363 | -11.2158
6.3973 | -34.1368 | -11.2259
6.4034 | -34.1371 | -11.2361
6.4095 | -34.1374 | -11.2466
6.4156 | -34.1378 | -11.2566
6.4217 | -34.1381 | -11.2669
6.4278 | -34.1384 | -11.2771
6.4339 | -34.1389 | -11.2874
6.44 | -34.1393 | -11.2973
6.4461 | -34.1396 | -11.308
6.4522 | -34.14 | -11.3177
6.4583 | -34.1403 | -11.3281
6.4644 | -34.1405 | -11.3387
6.4705 | -34.141 | -11.3491
6.4766 | -34.1413 | -11.3598
6.4827 | -34.1417 | -11.3705
6.4888 | -34.142 | -11.3806
6.4949 | -34.1424 | -11.3905
6.501 | -34.1427 | -11.4011
6.5071 | -34.1431 | -11.4115
6.5132 | -34.1435 | -11.4217
6.5194 | -34.1439 | -11.4319
6.5255 | -34.1443 | -11.4418
6.5316 | -34.1446 | -11.4522
6.5377 | -34.145 | -11.4627
6.5438 | -34.1453 | -11.4729
6.5499 | -34.1457 | -11.483
6.556 | -34.1461 | -11.4931
6.5621 | -34.1464 | -11.5035
6.5682 | -34.1467 | -11.5136
6.5743 | -34.1471 | -11.5246
6.5804 | -34.1474 | -11.5349
6.5865 | -34.1477 | -11.5454
6.5926 | -34.1481 | -11.5556
6.5987 | -34.1485 | -11.5656
6.6048 | -34.1488 | -11.5765
6.6109 | -34.1492 | -11.5867
6.617 | -34.1496 | -11.5978
6.6231 | -34.1498 | -11.6083
6.6292 | -34.1503 | -11.6185
6.6353 | -34.1506 | -11.6297
6.6414 | -34.1511 | -11.6403
6.6475 | -34.1514 | -11.6508
6.6536 | -34.1518 | -11.6608
6.6597 | -34.1522 | -11.6708
6.6659 | -34.1526 | -11.6809
6.672 | -34.153 | -11.6913
6.6781 | -34.1533 | -11.7014
6.6842 | -34.1537 | -11.7119
6.6903 | -34.1541 | -11.7221
6.6964 | -34.1544 | -11.7324
6.7025 | -34.1547 | -11.7427
6.7086 | -34.1551 | -11.7528
6.7147 | -34.1555 | -11.7632
6.7208 | -34.1558 | -11.774
6.7269 | -34.1561 | -11.784
6.733 | -34.1564 | -11.7935
6.7391 | -34.1569 | -11.804
6.7452 | -34.1573 | -11.8137
6.7513 | -34.1577 | -11.8237
6.7574 | -34.1581 | -11.8343
6.7635 | -34.1585 | -11.844
6.7696 | -34.1589 | -11.8544
6.7757 | -34.1593 | -11.865
6.7818 | -34.1597 | -11.8753
6.7879 | -34.16 | -11.8856
6.794 | -34.1604 | -11.8958
6.8001 | -34.1607 | -11.9058
6.8063 | -34.1611 | -11.9163
6.8124 | -34.1615 | -11.9266
6.8185 | -34.1618 | -11.9369
6.8246 | -34.1622 | -11.9471
6.8307 | -34.1625 | -11.9573
6.8368 | -34.1629 | -11.9678
6.8429 | -34.1633 | -11.978
6.849 | -34.1636 | -11.9882
6.8551 | -34.164 | -11.9985
6.8612 | -34.1644 | -12.0086
6.8673 | -34.1648 | -12.0186
6.8734 | -34.1651 | -12.0296
6.8795 | -34.1656 | -12.0401
6.8856 | -34.1659 | -12.0499
6.8917 | -34.1663 | -12.0596
6.8978 | -34.1666 | -12.0697
6.9039 | -34.167 | -12.0798
6.91 | -34.1674 | -12.0901
6.9161 | -34.1677 | -12.1002
6.9222 | -34.1681 | -12.1104
6.9283 | -34.1684 | -12.1203
6.9344 | -34.1688 | -12.1303
6.9405 | -34.1691 | -12.1404
6.9466 | -34.1695 | -12.1508
6.9528 | -34.1698 | -12.161
6.9589 | -34.1702 | -12.1711
6.965 | -34.1706 | -12.1813
6.9711 | -34.171 | -12.1915
6.9772 | -34.1714 | -12.2014
6.9833 | -34.1717 | -12.2118
6.9894 | -34.172 | -12.2224
6.9955 | -34.1724 | -12.2326
7.0016 | -34.1727 | -12.2431
7.0077 | -34.1731 | -12.2536
7.0138 | -34.1735 | -12.2641
7.0199 | -34.1738 | -12.2741
7.026 | -34.1741 | -12.2842
7.0321 | -34.1744 | -12.2944
7.0382 | -34.1748 | -12.3044
7.0443 | -34.1752 | -12.3149
7.0504 | -34.1755 | -12.3253
7.0565 | -34.1758 | -12.3357
7.0626 | -34.1762 | -12.3458
7.0687 | -34.1765 | -12.3556
7.0748 | -34.1769 | -12.3656
7.0809 | -34.1772 | -12.3759
7.087 | -34.1776 | -12.3861
7.0932 | -34.1779 | -12.3962
7.0993 | -34.1782 | -12.4064
7.1054 | -34.1786 | -12.4163
7.1115 | -34.1789 | -12.4264
7.1176 | -34.1792 | -12.4364
7.1237 | -34.1795 | -12.4465
7.1298 | -34.1798 | -12.4562
7.1359 | -34.1801 | -12.4661
7.142 | -34.1804 | -12.4763
7.1481 | -34.1807 | -12.4862
7.1542 | -34.181 | -12.4961
7.1603 | -34.1814 | -12.5061
7.1664 | -34.1818 | -12.5163
7.1725 | -34.1821 | -12.5262
7.1786 | -34.1825 | -12.5364
7.1847 | -34.1828 | -12.5463
7.1908 | -34.1831 | -12.5565
7.1969 | -34.1834 | -12.5665
7.203 | -34.1838 | -12.5766
7.2091 | -34.1841 | -12.5865
7.2152 | -34.1844 | -12.5965
7.2213 | -34.1847 | -12.6064
7.2274 | -34.1851 | -12.6164
7.2335 | -34.1854 | -12.6265
7.2397 | -34.1857 | -12.6365
7.2458 | -34.186 | -12.6461
7.2519 | -34.1864 | -12.6561
7.258 | -34.1867 | -12.6663
7.2641 | -34.187 | -12.6762
7.2702 | -34.1874 | -12.6865
7.2763 | -34.1877 | -12.6963
7.2824 | -34.1881 | -12.7062
7.2885 | -34.1885 | -12.7162
7.2946 | -34.1888 | -12.7264
7.3007 | -34.1891 | -12.7364
7.3068 | -34.1894 | -12.7467
7.3129 | -34.1897 | -12.7564
7.319 | -34.1901 | -12.7665
7.3251 | -34.1904 | -12.7762
7.3312 | -34.1908 | -12.7859
7.3373 | -34.1911 | -12.7956
7.3434 | -34.1915 | -12.8053
7.3495 | -34.1919 | -12.8148
7.3556 | -34.1922 | -12.8251
7.3617 | -34.1925 | -12.8355
7.3678 | -34.1928 | -12.8456
7.3739 | -34.1932 | -12.8551
7.3801 | -34.1935 | -12.865
7.3862 | -34.1938 | -12.8749
7.3923 | -34.1941 | -12.8847
7.3984 | -34.1945 | -12.8945
7.4045 | -34.1948 | -12.9043
7.4106 | -34.1952 | -12.9141
7.4167 | -34.1955 | -12.9242
7.4228 | -34.1958 | -12.934
7.4289 | -34.1961 | -12.9437
7.435 | -34.1965 | -12.9533
7.4411 | -34.1969 | -12.9629
7.4472 | -34.1972 | -12.9729
7.4533 | -34.1976 | -12.9828
7.4594 | -34.1979 | -12.993
7.4655 | -34.1983 | -13.0028
7.4716 | -34.1985 | -13.0129
7.4777 | -34.1989 | -13.0223
7.4838 | -34.1992 | -13.0324
7.4899 | -34.1996 | -13.0422
7.496 | -34.2 | -13.0522
7.5021 | -34.2003 | -13.0619
7.5082 | -34.2006 | -13.0718
7.5143 | -34.2009 | -13.0816
7.5204 | -34.2013 | -13.0911
7.5266 | -34.2016 | -13.101
7.5327 | -34.2019 | -13.1107
7.5388 | -34.2022 | -13.1205
7.5449 | -34.2025 | -13.1305
7.551 | -34.2028 | -13.1406
7.5571 | -34.2032 | -13.1503
7.5632 | -34.2035 | -13.1601
7.5693 | -34.2038 | -13.1701
7.5754 | -34.2042 | -13.1799
7.5815 | -34.2046 | -13.1893
7.5876 | -34.2048 | -13.1991
7.5937 | -34.2052 | -13.2088
7.5998 | -34.2055 | -13.2187
7.6059 | -34.2058 | -13.2284
7.612 | -34.2061 | -13.2382
7.6181 | -34.2064 | -13.248
7.6242 | -34.2068 | -13.2578
7.6303 | -34.2072 | -13.2673
7.6364 | -34.2075 | -13.2774
7.6425 | -34.2078 | -13.2875
7.6486 | -34.2081 | -13.2973
7.6547 | -34.2085 | -13.3072
7.6608 | -34.2089 | -13.317
7.667 | -34.2092 | -13.3263
7.6731 | -34.2095 | -13.3363
7.6792 | -34.2097 | -13.3462
7.6853 | -34.21 | -13.3558
7.6914 | -34.2103 | -13.3658
7.6975 | -34.2106 | -13.3757
7.7036 | -34.2109 | -13.3856
7.7097 | -34.2113 | -13.3953
7.7158 | -34.2116 | -13.405
7.7219 | -34.2119 | -13.4146
7.728 | -34.2122 | -13.4246
7.7341 | -34.2125 | -13.4344
7.7402 | -34.2129 | -13.4441
7.7463 | -34.2132 | -13.4539
7.7524 | -34.2135 | -13.464
7.7585 | -34.2138 | -13.4739
7.7646 | -34.2141 | -13.4834
7.7707 | -34.2143 | -13.493
7.7768 | -34.2146 | -13.5026
7.7829 | -34.2148 | -13.5125
7.789 | -34.2152 | -13.522
7.7951 | -34.2155 | -13.5322
7.8012 | -34.2158 | -13.5423
7.8073 | -34.2161 | -13.5519
7.8135 | -34.2164 | -13.5618
7.8196 | -34.2167 | -13.5715
7.8257 | -34.2171 | -13.5817
7.8318 | -34.2174 | -13.5919
7.8379 | -34.2177 | -13.6012
7.844 | -34.2179 | -13.6106
7.8501 | -34.2182 | -13.6207
7.8562 | -34.2184 | -13.6311
7.8623 | -34.2188 | -13.641
7.8684 | -34.2191 | -13.6505
7.8745 | -34.2194 | -13.6601
7.8806 | -34.2197 | -13.6693
7.8867 | -34.22 | -13.6789
7.8928 | -34.2204 | -13.6885
7.8989 | -34.2207 | -13.6979
7.905 | -34.2211 | -13.7081
7.9111 | -34.2214 | -13.7178
7.9172 | -34.2217 | -13.7273
7.9233 | -34.2219 | -13.7372
7.9294 | -34.2223 | -13.7471
7.9355 | -34.2226 | -13.7564
7.9416 | -34.2228 | -13.7661
7.9477 | -34.2231 | -13.7754
7.9539 | -34.2235 | -13.7848
7.96 | -34.2238 | -13.7948
7.9661 | -34.2241 | -13.8046
7.9722 | -34.2244 | -13.8145
7.9783 | -34.2247 | -13.8242
7.9844 | -34.2251 | -13.8344
7.9905 | -34.2254 | -13.844
7.9966 | -34.2256 | -13.8537
8.0027 | -34.2259 | -13.8638
8.0088 | -34.2262 | -13.8731
8.0149 | -34.2265 | -13.8826
8.021 | -34.2268 | -13.8923
8.0271 | -34.2271 | -13.9025
8.0332 | -34.2274 | -13.9121
8.0393 | -34.2278 | -13.9221
8.0454 | -34.228 | -13.932
8.0515 | -34.2284 | -13.9416
8.0576 | -34.2286 | -13.9517
8.0637 | -34.2289 | -13.9609
8.0698 | -34.2292 | -13.9703
8.0759 | -34.2295 | -13.9797
8.082 | -34.2299 | -13.9897
8.0881 | -34.2301 | -13.9993
8.0942 | -34.2304 | -14.0094
8.1004 | -34.2307 | -14.0195
8.1065 | -34.2311 | -14.0292
8.1126 | -34.2314 | -14.0388
8.1187 | -34.2317 | -14.0486
8.1248 | -34.232 | -14.0586
8.1309 | -34.2323 | -14.0683
8.137 | -34.2326 | -14.0779
8.1431 | -34.2329 | -14.0878
8.1492 | -34.2332 | -14.0976
8.1553 | -34.2336 | -14.1074
8.1614 | -34.2339 | -14.1174
8.1675 | -34.2341 | -14.127
8.1736 | -34.2344 | -14.137
8.1797 | -34.2347 | -14.1469
8.1858 | -34.235 | -14.1566
8.1919 | -34.2353 | -14.1663
8.198 | -34.2356 | -14.1757
8.2041 | -34.2359 | -14.1858
8.2102 | -34.2363 | -14.1955
8.2163 | -34.2366 | -14.2051
8.2224 | -34.2369 | -14.2151
8.2285 | -34.2372 | -14.2246
8.2346 | -34.2375 | -14.2343
8.2408 | -34.2378 | -14.2442
8.2469 | -34.2381 | -14.2537
8.253 | -34.2384 | -14.2634
8.2591 | -34.2387 | -14.2729
8.2652 | -34.2391 | -14.2824
8.2713 | -34.2394 | -14.2922
8.2774 | -34.2397 | -14.302
8.2835 | -34.2401 | -14.3121
8.2896 | -34.2403 | -14.3214
8.2957 | -34.2407 | -14.3311
8.3018 | -34.241 | -14.3408
8.3079 | -34.2413 | -14.3505
8.314 | -34.2416 | -14.3601
8.3201 | -34.242 | -14.3697
8.3262 | -34.2422 | -14.3794
8.3323 | -34.2426 | -14.3891
8.3384 | -34.2429 | -14.3992
8.3445 | -34.2432 | -14.4085
8.3506 | -34.2435 | -14.4184
8.3567 | -34.2438 | -14.4283
8.3628 | -34.2441 | -14.438
8.3689 | -34.2444 | -14.4478
8.375 | -34.2447 | -14.4575
8.3812 | -34.245 | -14.4672
8.3873 | -34.2454 | -14.477
8.3934 | -34.2457 | -14.4867
8.3995 | -34.246 | -14.4965
8.4056 | -34.2463 | -14.5063
8.4117 | -34.2466 | -14.5163
8.4178 | -34.2469 | -14.5256
8.4239 | -34.2473 | -14.5351
8.43 | -34.2476 | -14.5449
8.4361 | -34.2479 | -14.555
8.4422 | -34.2482 | -14.5647
8.4483 | -34.2485 | -14.5744
8.4544 | -34.2488 | -14.5843
8.4605 | -34.2491 | -14.594
8.4666 | -34.2494 | -14.6039
8.4727 | -34.2497 | -14.6136
8.4788 | -34.25 | -14.6231
8.4849 | -34.2503 | -14.6331
8.491 | -34.2506 | -14.6427
8.4971 | -34.251 | -14.6525
8.5032 | -34.2513 | -14.6623
8.5093 | -34.2516 | -14.672
8.5154 | -34.2519 | -14.6815
8.5215 | -34.2522 | -14.6913
8.5277 | -34.2525 | -14.7011
8.5338 | -34.2528 | -14.7105
8.5399 | -34.2531 | -14.7205
8.546 | -34.2535 | -14.73
8.5521 | -34.2538 | -14.7398
8.5582 | -34.2541 | -14.7496
8.5643 | -34.2544 | -14.7592
8.5704 | -34.2547 | -14.7686
8.5765 | -34.255 | -14.7783
8.5826 | -34.2554 | -14.788
8.5887 | -34.2557 | -14.7976
8.5948 | -34.256 | -14.807
8.6009 | -34.2563 | -14.8166
8.607 | -34.2566 | -14.8266
8.6131 | -34.2569 | -14.8366
8.6192 | -34.2572 | -14.8459
8.6253 | -34.2575 | -14.8558
8.6314 | -34.2579 | -14.8655
8.6375 | -34.2582 | -14.8754
8.6436 | -34.2586 | -14.8851
8.6497 | -34.2589 | -14.8949
8.6558 | -34.2592 | -14.9044
8.6619 | -34.2595 | -14.9144
8.6681 | -34.2598 | -14.9239
8.6742 | -34.2601 | -14.9339
8.6803 | -34.2605 | -14.9439
8.6864 | -34.2608 | -14.9539
8.6925 | -34.2612 | -14.9636
8.6986 | -34.2615 | -14.9736
8.7047 | -34.2618 | -14.9834
8.7108 | -34.2621 | -14.9933
8.7169 | -34.2624 | -15.0027
8.723 | -34.2628 | -15.0125
8.7291 | -34.2631 | -15.0221
8.7352 | -34.2635 | -15.0321
8.7413 | -34.2638 | -15.0418
8.7474 | -34.2641 | -15.0512
8.7535 | -34.2645 | -15.0611
8.7596 | -34.2648 | -15.0713
8.7657 | -34.2651 | -15.081
8.7718 | -34.2654 | -15.0905
8.7779 | -34.2658 | -15.1008
8.784 | -34.266 | -15.1104
8.7901 | -34.2663 | -15.1205
8.7962 | -34.2667 | -15.1304
8.8023 | -34.2671 | -15.1398
8.8084 | -34.2674 | -15.1497
8.8146 | -34.2677 | -15.1593
8.8207 | -34.268 | -15.1691
8.8268 | -34.2684 | -15.1788
8.8329 | -34.2687 | -15.1887
8.839 | -34.2691 | -15.1985
8.8451 | -34.2694 | -15.2081
8.8512 | -34.2698 | -15.2179
8.8573 | -34.2701 | -15.2274
8.8634 | -34.2705 | -15.2371
8.8695 | -34.2708 | -15.247
8.8756 | -34.2711 | -15.2565
8.8817 | -34.2715 | -15.2665
8.8878 | -34.2717 | -15.2765
8.8939 | -34.272 | -15.2859
8.9 | -34.2724 | -15.2954
8.9061 | -34.2728 | -15.3052
8.9122 | -34.2731 | -15.3148
8.9183 | -34.2734 | -15.3244
8.9244 | -34.2737 | -15.334
8.9305 | -34.2741 | -15.3438
8.9366 | -34.2745 | -15.3534
8.9427 | -34.2749 | -15.3632
8.9488 | -34.2752 | -15.3728
8.955 | -34.2755 | -15.3825
8.9611 | -34.2759 | -15.3924
8.9672 | -34.2763 | -15.4025
8.9733 | -34.2766 | -15.4122
8.9794 | -34.2769 | -15.422
8.9855 | -34.2773 | -15.4316
8.9916 | -34.2776 | -15.4412
8.9977 | -34.278 | -15.4509
9.0038 | -34.2784 | -15.4607
9.0099 | -34.2787 | -15.4706
9.016 | -34.2791 | -15.4803
9.0221 | -34.2794 | -15.4903
9.0282 | -34.2798 | -15.5002
9.0343 | -34.2801 | -15.5101
9.0404 | -34.2806 | -15.5198
9.0465 | -34.2809 | -15.5299
9.0526 | -34.2813 | -15.5396
9.0587 | -34.2817 | -15.5492
9.0648 | -34.282 | -15.5589
9.0709 | -34.2824 | -15.5684
9.077 | -34.2827 | -15.5786
9.0831 | -34.2831 | -15.5884
9.0892 | -34.2834 | -15.598
9.0953 | -34.2837 | -15.6079
9.1015 | -34.284 | -15.6176
9.1076 | -34.2844 | -15.6275
9.1137 | -34.2848 | -15.6372
9.1198 | -34.2852 | -15.6466
9.1259 | -34.2856 | -15.6564
9.132 | -34.286 | -15.6664
9.1381 | -34.2864 | -15.676
9.1442 | -34.2868 | -15.6854
9.1503 | -34.2871 | -15.6952
9.1564 | -34.2875 | -15.705
9.1625 | -34.2878 | -15.7148
9.1686 | -34.2882 | -15.7245
9.1747 | -34.2886 | -15.7343
9.1808 | -34.2889 | -15.7443
9.1869 | -34.2893 | -15.7543
9.193 | -34.2896 | -15.7642
9.1991 | -34.2899 | -15.7737
9.2052 | -34.2903 | -15.7833
9.2113 | -34.2907 | -15.7933
9.2174 | -34.2911 | -15.8036
9.2235 | -34.2915 | -15.8131
9.2296 | -34.2919 | -15.823
9.2357 | -34.2923 | -15.8327
9.2419 | -34.2926 | -15.8425
9.248 | -34.293 | -15.8524
9.2541 | -34.2934 | -15.8623
9.2602 | -34.2937 | -15.8721
9.2663 | -34.2941 | -15.8817
9.2724 | -34.2945 | -15.8915
9.2785 | -34.2948 | -15.9012
9.2846 | -34.2953 | -15.911
9.2907 | -34.2956 | -15.9207
9.2968 | -34.2959 | -15.9306
9.3029 | -34.2963 | -15.9405
9.309 | -34.2967 | -15.9502
9.3151 | -34.2971 | -15.9599
9.3212 | -34.2975 | -15.9698
9.3273 | -34.2979 | -15.9794
9.3334 | -34.2983 | -15.989
9.3395 | -34.2988 | -15.9986
9.3456 | -34.2992 | -16.0081
9.3517 | -34.2996 | -16.0175
9.3578 | -34.2999 | -16.0271
9.3639 | -34.3004 | -16.037
9.37 | -34.3007 | -16.0467
9.3761 | -34.3011 | -16.0563
9.3822 | -34.3015 | -16.0659
9.3884 | -34.3019 | -16.0757
9.3945 | -34.3023 | -16.0855
9.4006 | -34.3027 | -16.0953
9.4067 | -34.303 | -16.1052
9.4128 | -34.3034 | -16.1148
9.4189 | -34.3038 | -16.1247
9.425 | -34.3042 | -16.1343
9.4311 | -34.3046 | -16.144
9.4372 | -34.305 | -16.1538
9.4433 | -34.3054 | -16.1636
9.4494 | -34.3058 | -16.1732
9.4555 | -34.3062 | -16.1829
9.4616 | -34.3066 | -16.1926
9.4677 | -34.3071 | -16.2022
9.4738 | -34.3075 | -16.2123
9.4799 | -34.3078 | -16.222
9.486 | -34.3081 | -16.2321
9.4921 | -34.3085 | -16.2422
9.4982 | -34.309 | -16.2515
9.5043 | -34.3093 | -16.2609
9.5104 | -34.3098 | -16.2704
9.5165 | -34.3102 | -16.2801
9.5226 | -34.3106 | -16.29
9.5288 | -34.311 | -16.2997
9.5349 | -34.3114 | -16.3094
9.541 | -34.3118 | -16.3192
9.5471 | -34.3121 | -16.329
9.5532 | -34.3125 | -16.3386
9.5593 | -34.313 | -16.3482
9.5654 | -34.3134 | -16.3579
9.5715 | -34.3138 | -16.3679
9.5776 | -34.3142 | -16.3779
9.5837 | -34.3145 | -16.3873
9.5898 | -34.3149 | -16.3972
9.5959 | -34.3153 | -16.4068
9.602 | -34.3156 | -16.4166
9.6081 | -34.316 | -16.4264
9.6142 | -34.3165 | -16.4359
9.6203 | -34.3168 | -16.4457
9.6264 | -34.3172 | -16.455
9.6325 | -34.3175 | -16.4649
9.6386 | -34.3179 | -16.4745
9.6447 | -34.3183 | -16.4841
9.6508 | -34.3187 | -16.4938
9.6569 | -34.3191 | -16.5035
9.663 | -34.3196 | -16.5132
9.6691 | -34.32 | -16.5229
9.6753 | -34.3204 | -16.5326
9.6814 | -34.3208 | -16.542
9.6875 | -34.3213 | -16.5515
9.6936 | -34.3217 | -16.561
9.6997 | -34.3221 | -16.5708
9.7058 | -34.3225 | -16.5805
9.7119 | -34.3229 | -16.5902
9.718 | -34.3232 | -16.6002
9.7241 | -34.3236 | -16.6099
9.7302 | -34.324 | -16.6196
9.7363 | -34.3244 | -16.629
9.7424 | -34.3249 | -16.6386
9.7485 | -34.3252 | -16.6481
9.7546 | -34.3257 | -16.6575
9.7607 | -34.326 | -16.6674
9.7668 | -34.3265 | -16.6775
9.7729 | -34.3269 | -16.6868
9.779 | -34.3274 | -16.6965
9.7851 | -34.3277 | -16.7062
9.7912 | -34.3281 | -16.7164
9.7973 | -34.3285 | -16.7263
9.8034 | -34.3289 | -16.7362
9.8095 | -34.3293 | -16.7457
9.8157 | -34.3297 | -16.7552
9.8218 | -34.3301 | -16.7645
9.8279 | -34.3305 | -16.7742
9.834 | -34.3309 | -16.7838
9.8401 | -34.3313 | -16.7933
9.8462 | -34.3318 | -16.8025
9.8523 | -34.3322 | -16.8121
9.8584 | -34.3326 | -16.822
9.8645 | -34.333 | -16.8316
9.8706 | -34.3333 | -16.8411
9.8767 | -34.3337 | -16.8504
9.8828 | -34.3342 | -16.8596
9.8889 | -34.3346 | -16.8695
9.895 | -34.335 | -16.8788
9.9011 | -34.3354 | -16.888
9.9072 | -34.3358 | -16.8973
9.9133 | -34.3362 | -16.9068
9.9194 | -34.3366 | -16.9161
9.9255 | -34.337 | -16.9258
9.9316 | -34.3375 | -16.9351
9.9377 | -34.3379 | -16.9444
9.9438 | -34.3383 | -16.9539
9.9499 | -34.3387 | -16.9636
9.956 | -34.3391 | -16.9729
9.9622 | -34.3395 | -16.9824
9.9683 | -34.3399 | -16.9921
9.9744 | -34.3403 | -17.0015
9.9805 | -34.3407 | -17.0111
9.9866 | -34.3411 | -17.0203
9.9927 | -34.3416 | -17.0299
9.9988 | -34.342 | -17.0393
10.0049 | -34.3424 | -17.0488
10.011 | -34.3428 | -17.0582
10.0171 | -34.3432 | -17.0678
10.0232 | -34.3436 | -17.0771
10.0293 | -34.3441 | -17.0864
10.0354 | -34.3446 | -17.0956
10.0415 | -34.345 | -17.1051
10.0476 | -34.3453 | -17.1149
10.0537 | -34.3457 | -17.1246
10.0598 | -34.3461 | -17.1341
10.0659 | -34.3465 | -17.1436
10.072 | -34.347 | -17.1529
10.0781 | -34.3474 | -17.1623
10.0842 | -34.3478 | -17.1719
10.0903 | -34.3482 | -17.1817
10.0964 | -34.3487 | -17.1909
10.1026 | -34.3491 | -17.2004
10.1087 | -34.3496 | -17.2099
10.1148 | -34.35 | -17.2194
10.1209 | -34.3504 | -17.2287
10.127 | -34.3508 | -17.2379
10.1331 | -34.3513 | -17.2472
10.1392 | -34.3517 | -17.2566
10.1453 | -34.3521 | -17.2661
10.1514 | -34.3525 | -17.2753
10.1575 | -34.3529 | -17.2847
10.1636 | -34.3534 | -17.2934
10.1697 | -34.3538 | -17.3028
10.1758 | -34.3542 | -17.312
10.1819 | -34.3546 | -17.3215
10.188 | -34.355 | -17.3306
10.1941 | -34.3554 | -17.3401
10.2002 | -34.3557 | -17.3487
10.2063 | -34.3562 | -17.3584
10.2124 | -34.3566 | -17.3679
10.2185 | -34.357 | -17.3774
10.2246 | -34.3574 | -17.3867
10.2307 | -34.3578 | -17.3963
10.2368 | -34.3582 | -17.4049
10.2429 | -34.3587 | -17.414
10.2491 | -34.3591 | -17.4232
10.2552 | -34.3595 | -17.433
10.2613 | -34.36 | -17.4426
10.2674 | -34.3603 | -17.4519
10.2735 | -34.3607 | -17.4615
10.2796 | -34.3611 | -17.471
10.2857 | -34.3614 | -17.4804
10.2918 | -34.3619 | -17.49
10.2979 | -34.3623 | -17.4994
10.304 | -34.3628 | -17.5089
10.3101 | -34.3632 | -17.5183
10.3162 | -34.3636 | -17.5277
10.3223 | -34.364 | -17.537
10.3284 | -34.3645 | -17.5462
10.3345 | -34.3649 | -17.5555
10.3406 | -34.3653 | -17.5647
10.3467 | -34.3657 | -17.574
10.3528 | -34.3661 | -17.583
10.3589 | -34.3664 | -17.5925
10.365 | -34.3669 | -17.6021
10.3711 | -34.3672 | -17.6117
10.3772 | -34.3676 | -17.6211
10.3833 | -34.368 | -17.6303
10.3895 | -34.3684 | -17.6395
10.3956 | -34.3689 | -17.6485
10.4017 | -34.3692 | -17.6577
10.4078 | -34.3697 | -17.6671
10.4139 | -34.3701 | -17.6763
10.42 | -34.3705 | -17.6857
10.4261 | -34.371 | -17.6949
10.4322 | -34.3714 | -17.704
10.4383 | -34.3718 | -17.7133
10.4444 | -34.3722 | -17.7221
10.4505 | -34.3726 | -17.7314
10.4566 | -34.373 | -17.7405
10.4627 | -34.3734 | -17.7498
10.4688 | -34.3738 | -17.7586
10.4749 | -34.3742 | -17.7677
10.481 | -34.3746 | -17.777
10.4871 | -34.3751 | -17.7863
10.4932 | -34.3755 | -17.7955
10.4993 | -34.376 | -17.8046
10.5054 | -34.3764 | -17.814
10.5115 | -34.3768 | -17.8232
10.5176 | -34.3772 | -17.8323
10.5237 | -34.3776 | -17.8415
10.5298 | -34.378 | -17.8505
10.536 | -34.3785 | -17.86
10.5421 | -34.3789 | -17.8687
10.5482 | -34.3792 | -17.8783
10.5543 | -34.3797 | -17.8873
10.5604 | -34.3801 | -17.8962
10.5665 | -34.3805 | -17.9052
10.5726 | -34.3809 | -17.9142
10.5787 | -34.3814 | -17.9236
10.5848 | -34.3818 | -17.9325
10.5909 | -34.3822 | -17.9418
10.597 | -34.3826 | -17.9508
10.6031 | -34.383 | -17.9602
10.6092 | -34.3834 | -17.9696
10.6153 | -34.3838 | -17.979
10.6214 | -34.3842 | -17.9879
10.6275 | -34.3847 | -17.9971
10.6336 | -34.3851 | -18.0067
10.6397 | -34.3855 | -18.0158
10.6458 | -34.3859 | -18.0249
10.6519 | -34.3862 | -18.034
10.658 | -34.3866 | -18.0428
10.6641 | -34.3871 | -18.052
10.6702 | -34.3875 | -18.0614
10.6764 | -34.3879 | -18.0706
10.6825 | -34.3884 | -18.08
10.6886 | -34.3888 | -18.0892
10.6947 | -34.3892 | -18.0983
10.7008 | -34.3897 | -18.1073
10.7069 | -34.3901 | -18.1166
10.713 | -34.3905 | -18.1256
10.7191 | -34.3909 | -18.1348
10.7252 | -34.3914 | -18.144
10.7313 | -34.3918 | -18.1532
10.7374 | -34.3922 | -18.1621
10.7435 | -34.3926 | -18.1713
10.7496 | -34.393 | -18.1803
10.7557 | -34.3934 | -18.1893
10.7618 | -34.3939 | -18.1984
10.7679 | -34.3943 | -18.2074
10.774 | -34.3946 | -18.2166
10.7801 | -34.3951 | -18.2263
10.7862 | -34.3955 | -18.2353
10.7923 | -34.3959 | -18.2445
10.7984 | -34.3963 | -18.2535
10.8045 | -34.3967 | -18.2629
10.8106 | -34.3972 | -18.272
10.8168 | -34.3976 | -18.2812
10.8229 | -34.398 | -18.2903
10.829 | -34.3984 | -18.2995
10.8351 | -34.3988 | -18.3086
10.8412 | -34.3992 | -18.3176
10.8473 | -34.3996 | -18.3263
10.8534 | -34.4001 | -18.3356
10.8595 | -34.4005 | -18.3446
10.8656 | -34.401 | -18.3538
10.8717 | -34.4014 | -18.3625
10.8778 | -34.4018 | -18.3713
10.8839 | -34.4022 | -18.3804
10.89 | -34.4026 | -18.3894
10.8961 | -34.4031 | -18.3986
10.9022 | -34.4035 | -18.4075
10.9083 | -34.4039 | -18.4168
10.9144 | -34.4044 | -18.4257
10.9205 | -34.4048 | -18.4348
10.9266 | -34.4053 | -18.4442
10.9327 | -34.4058 | -18.4533
10.9388 | -34.4062 | -18.4623
10.9449 | -34.4065 | -18.4715
10.951 | -34.4069 | -18.4805
10.9571 | -34.4074 | -18.4897
10.9633 | -34.4078 | -18.4989
10.9694 | -34.4082 | -18.5082
10.9755 | -34.4086 | -18.5173
10.9816 | -34.409 | -18.5265
10.9877 | -34.4095 | -18.5355
10.9938 | -34.4099 | -18.5448
10.9999 | -34.4103 | -18.5538
11.006 | -34.4108 | -18.5628
11.0121 | -34.4112 | -18.5723
11.0182 | -34.4116 | -18.581
11.0243 | -34.4121 | -18.5902
11.0304 | -34.4124 | -18.5992
11.0365 | -34.4129 | -18.6083
11.0426 | -34.4133 | -18.6171
11.0487 | -34.4137 | -18.6262
11.0548 | -34.4142 | -18.6353
11.0609 | -34.4146 | -18.6446
11.067 | -34.415 | -18.6539
11.0731 | -34.4154 | -18.663
11.0792 | -34.4158 | -18.6721
11.0853 | -34.4162 | -18.6808
11.0914 | -34.4167 | -18.6899
11.0975 | -34.4171 | -18.6992
11.1037 | -34.4175 | -18.708
11.1098 | -34.4179 | -18.7174
11.1159 | -34.4184 | -18.7266
11.122 | -34.4188 | -18.7359
11.1281 | -34.4192 | -18.7447
11.1342 | -34.4196 | -18.7534
11.1403 | -34.42 | -18.7627
11.1464 | -34.4205 | -18.7717
11.1525 | -34.4209 | -18.7808
11.1586 | -34.4213 | -18.7903
11.1647 | -34.4218 | -18.7991
11.1708 | -34.4222 | -18.808
11.1769 | -34.4225 | -18.8165
11.183 | -34.423 | -18.8255
11.1891 | -34.4234 | -18.8349
11.1952 | -34.4239 | -18.8443
11.2013 | -34.4243 | -18.8533
11.2074 | -34.4247 | -18.8621
11.2135 | -34.425 | -18.8712
11.2196 | -34.4254 | -18.8802
11.2257 | -34.4258 | -18.8892
11.2318 | -34.4261 | -18.8986
11.2379 | -34.4265 | -18.9075
11.244 | -34.4269 | -18.9166
11.2502 | -34.4273 | -18.9257
11.2563 | -34.4278 | -18.935
11.2624 | -34.4282 | -18.9441
11.2685 | -34.4286 | -18.953
11.2746 | -34.429 | -18.9618
11.2807 | -34.4294 | -18.9712
11.2868 | -34.4297 | -18.9804
11.2929 | -34.4302 | -18.9893
11.299 | -34.4306 | -18.9987
11.3051 | -34.431 | -19.0077
11.3112 | -34.4314 | -19.0168
11.3173 | -34.4319 | -19.0258
11.3234 | -34.4323 | -19.0351
11.3295 | -34.4327 | -19.044
11.3356 | -34.4331 | -19.0531
11.3417 | -34.4335 | -19.0623
11.3478 | -34.4339 | -19.0714
11.3539 | -34.4343 | -19.0804
11.36 | -34.4347 | -19.0895
11.3661 | -34.4352 | -19.0986
11.3722 | -34.4356 | -19.1076
11.3783 | -34.436 | -19.1167
11.3844 | -34.4364 | -19.1258
11.3906 | -34.4368 | -19.1353
11.3967 | -34.4372 | -19.1443
11.4028 | -34.4376 | -19.1537
11.4089 | -34.438 | -19.1625
11.415 | -34.4384 | -19.1716
11.4211 | -34.4388 | -19.1806
11.4272 | -34.4392 | -19.1897
11.4333 | -34.4396 | -19.1992
11.4394 | -34.4401 | -19.2085
11.4455 | -34.4404 | -19.2178
11.4516 | -34.4408 | -19.2269
11.4577 | -34.4413 | -19.2361
11.4638 | -34.4418 | -19.2447
11.4699 | -34.4421 | -19.2542
11.476 | -34.4425 | -19.2639
11.4821 | -34.443 | -19.2726
11.4882 | -34.4434 | -19.2818
11.4943 | -34.4438 | -19.2908
11.5004 | -34.4443 | -19.3
11.5065 | -34.4447 | -19.309
11.5126 | -34.4451 | -19.3178
11.5187 | -34.4456 | -19.3268
11.5248 | -34.446 | -19.3358
11.5309 | -34.4464 | -19.345
11.5371 | -34.4468 | -19.354
11.5432 | -34.4471 | -19.3632
11.5493 | -34.4475 | -19.3725
11.5554 | -34.4479 | -19.3813
11.5615 | -34.4484 | -19.3906
11.5676 | -34.4488 | -19.3998
11.5737 | -34.4492 | -19.4087
11.5798 | -34.4496 | -19.4175
11.5859 | -34.4501 | -19.4264
11.592 | -34.4505 | -19.4355
11.5981 | -34.4509 | -19.4445
11.6042 | -34.4513 | -19.4537
11.6103 | -34.4517 | -19.4625
11.6164 | -34.4522 | -19.4714
11.6225 | -34.4526 | -19.4804
11.6286 | -34.453 | -19.4895
11.6347 | -34.4534 | -19.4989
11.6408 | -34.4539 | -19.5079
11.6469 | -34.4543 | -19.5172
11.653 | -34.4547 | -19.5261
11.6591 | -34.4552 | -19.5352
11.6652 | -34.4556 | -19.5444
11.6713 | -34.4559 | -19.5534
11.6775 | -34.4564 | -19.5624
11.6836 | -34.4568 | -19.5716
11.6897 | -34.4573 | -19.5806
11.6958 | -34.4577 | -19.5899
11.7019 | -34.4581 | -19.5992
11.708 | -34.4585 | -19.6085
11.7141 | -34.459 | -19.6175
11.7202 | -34.4594 | -19.6264
11.7263 | -34.4598 | -19.6356
11.7324 | -34.4603 | -19.6448
11.7385 | -34.4607 | -19.6541
11.7446 | -34.4611 | -19.6634
11.7507 | -34.4615 | -19.6723
11.7568 | -34.4619 | -19.6814
11.7629 | -34.4624 | -19.6905
11.769 | -34.4628 | -19.6995
11.7751 | -34.4632 | -19.7082
11.7812 | -34.4637 | -19.7172
11.7873 | -34.4641 | -19.7265
11.7934 | -34.4646 | -19.7357
11.7995 | -34.465 | -19.7443
11.8056 | -34.4654 | -19.7538
11.8117 | -34.4657 | -19.763
11.8178 | -34.4662 | -19.7721
11.824 | -34.4667 | -19.781
11.8301 | -34.4671 | -19.7896
11.8362 | -34.4675 | -19.7984
11.8423 | -34.4679 | -19.8075
11.8484 | -34.4684 | -19.8166
11.8545 | -34.4688 | -19.8262
11.8606 | -34.4692 | -19.8352
11.8667 | -34.4697 | -19.8444
11.8728 | -34.4701 | -19.8534
11.8789 | -34.4705 | -19.8628
11.885 | -34.4709 | -19.8717
11.8911 | -34.4714 | -19.881
11.8972 | -34.4718 | -19.8898
11.9033 | -34.4722 | -19.8985
11.9094 | -34.4726 | -19.9072
11.9155 | -34.4729 | -19.9163
11.9216 | -34.4734 | -19.9253
11.9277 | -34.4739 | -19.9341
11.9338 | -34.4743 | -19.9433
11.9399 | -34.4747 | -19.9526
11.946 | -34.4751 | -19.9617
11.9521 | -34.4756 | -19.9706
11.9582 | -34.476 | -19.9797
11.9644 | -34.4765 | -19.9886
11.9705 | -34.4769 | -19.9978
11.9766 | -34.4773 | -20.0067
11.9827 | -34.4777 | -20.0157
11.9888 | -34.4781 | -20.0245
11.9949 | -34.4785 | -20.0335
12.001 | -34.479 | -20.0426
12.0071 | -34.4794 | -20.0514
12.0132 | -34.4798 | -20.0601
12.0193 | -34.4803 | -20.0694
12.0254 | -34.4807 | -20.0782
12.0315 | -34.4812 | -20.0872
12.0376 | -34.4816 | -20.0959
12.0437 | -34.482 | -20.1046
12.0498 | -34.4824 | -20.1139
12.0559 | -34.4828 | -20.1229
12.062 | -34.4833 | -20.1318
12.0681 | -34.4837 | -20.1407
12.0742 | -34.4841 | -20.1499
12.0803 | -34.4846 | -20.1588
12.0864 | -34.485 | -20.1678
12.0925 | -34.4855 | -20.1769
12.0986 | -34.4859 | -20.1856
12.1047 | -34.4863 | -20.1945
12.1109 | -34.4868 | -20.2035
12.117 | -34.4873 | -20.2123
12.1231 | -34.4877 | -20.2213
12.1292 | -34.4881 | -20.2305
12.1353 | -34.4886 | -20.2392
12.1414 | -34.4891 | -20.2484
12.1475 | -34.4895 | -20.2574
12.1536 | -34.4899 | -20.2664
12.1597 | -34.4903 | -20.2755
12.1658 | -34.4907 | -20.2845
12.1719 | -34.4913 | -20.2937
12.178 | -34.4917 | -20.3027
12.1841 | -34.4922 | -20.312
12.1902 | -34.4926 | -20.3211
12.1963 | -34.4931 | -20.3301
12.2024 | -34.4935 | -20.339
12.2085 | -34.4939 | -20.3481
12.2146 | -34.4944 | -20.3569
12.2207 | -34.4948 | -20.3658
12.2268 | -34.4952 | -20.3748
12.2329 | -34.4957 | -20.3837
12.239 | -34.4961 | -20.3924
12.2451 | -34.4965 | -20.4011
12.2513 | -34.497 | -20.4101
12.2574 | -34.4974 | -20.4193
12.2635 | -34.4979 | -20.4279
12.2696 | -34.4983 | -20.4366
12.2757 | -34.4987 | -20.4452
12.2818 | -34.4992 | -20.454
12.2879 | -34.4996 | -20.4628
12.294 | -34.5 | -20.4715
12.3001 | -34.5005 | -20.4807
12.3062 | -34.5009 | -20.4896
12.3123 | -34.5014 | -20.4986
12.3184 | -34.5018 | -20.5074
12.3245 | -34.5023 | -20.5164
12.3306 | -34.5027 | -20.5251
12.3367 | -34.5032 | -20.5338
12.3428 | -34.5036 | -20.5425
12.3489 | -34.5041 | -20.5508
12.355 | -34.5045 | -20.5597
12.3611 | -34.5049 | -20.5687
12.3672 | -34.5054 | -20.5774
12.3733 | -34.5058 | -20.5867
12.3794 | -34.5062 | -20.5951
12.3855 | -34.5067 | -20.6038
12.3916 | -34.5071 | -20.6129
12.3978 | -34.5075 | -20.622
12.4039 | -34.5079 | -20.6309
12.41 | -34.5084 | -20.64
12.4161 | -34.5089 | -20.6486
12.4222 | -34.5093 | -20.6573
12.4283 | -34.5098 | -20.6661
12.4344 | -34.5102 | -20.675
12.4405 | -34.5107 | -20.6839
12.4466 | -34.5111 | -20.6926
12.4527 | -34.5116 | -20.7015
12.4588 | -34.512 | -20.7103
12.4649 | -34.5125 | -20.7194
12.471 | -34.5129 | -20.7287
12.4771 | -34.5134 | -20.737
12.4832 | -34.5138 | -20.7458
12.4893 | -34.5142 | -20.7544
12.4954 | -34.5147 | -20.7631
12.5015 | -34.5152 | -20.7717
12.5076 | -34.5156 | -20.7806
12.5137 | -34.5161 | -20.7899
12.5198 | -34.5165 | -20.7984
12.5259 | -34.517 | -20.8071
12.532 | -34.5175 | -20.8161
12.5382 | -34.5179 | -20.8247
12.5443 | -34.5183 | -20.8336
12.5504 | -34.5188 | -20.8424
12.5565 | -34.5193 | -20.8511
12.5626 | -34.5198 | -20.8601
12.5687 | -34.5202 | -20.8691
12.5748 | -34.5207 | -20.8779
12.5809 | -34.5211 | -20.8864
12.587 | -34.5215 | -20.8954
12.5931 | -34.522 | -20.9042
12.5992 | -34.5224 | -20.9133
12.6053 | -34.5229 | -20.9219
12.6114 | -34.5234 | -20.9307
12.6175 | -34.5238 | -20.9396
12.6236 | -34.5242 | -20.9482
12.6297 | -34.5247 | -20.9567
12.6358 | -34.5252 | -20.9653
12.6419 | -34.5257 | -20.9743
12.648 | -34.5261 | -20.9834
12.6541 | -34.5266 | -20.9921
12.6602 | -34.5271 | -21.0005
12.6663 | -34.5275 | -21.0092
12.6724 | -34.5279 | -21.018
12.6785 | -34.5284 | -21.027
12.6847 | -34.5288 | -21.0357
12.6908 | -34.5293 | -21.0444
12.6969 | -34.5298 | -21.053
12.703 | -34.5303 | -21.0619
12.7091 | -34.5307 | -21.0707
12.7152 | -34.5312 | -21.0794
12.7213 | -34.5316 | -21.0883
12.7274 | -34.5321 | -21.0971
12.7335 | -34.5325 | -21.1052
12.7396 | -34.533 | -21.1136
12.7457 | -34.5334 | -21.1223
12.7518 | -34.5339 | -21.131
12.7579 | -34.5343 | -21.14
12.764 | -34.5347 | -21.1488
12.7701 | -34.5351 | -21.157
12.7762 | -34.5355 | -21.1658
12.7823 | -34.536 | -21.1745
12.7884 | -34.5364 | -21.1831
12.7945 | -34.5369 | -21.1915
12.8006 | -34.5373 | -21.1995
12.8067 | -34.5378 | -21.2085
12.8128 | -34.5382 | -21.2168
12.8189 | -34.5386 | -21.2253
12.8251 | -34.539 | -21.2337
12.8312 | -34.5394 | -21.2424
12.8373 | -34.5399 | -21.251
12.8434 | -34.5404 | -21.2596
12.8495 | -34.5408 | -21.2681
12.8556 | -34.5412 | -21.2767
12.8617 | -34.5417 | -21.2851
12.8678 | -34.5421 | -21.2937
12.8739 | -34.5426 | -21.3019
12.88 | -34.543 | -21.3107
12.8861 | -34.5434 | -21.3195
12.8922 | -34.5438 | -21.328
12.8983 | -34.5443 | -21.3362
12.9044 | -34.5448 | -21.3448
12.9105 | -34.5453 | -21.3532
12.9166 | -34.5458 | -21.3622
12.9227 | -34.5462 | -21.3713
12.9288 | -34.5467 | -21.3801
12.9349 | -34.5472 | -21.389
12.941 | -34.5477 | -21.3975
12.9471 | -34.5482 | -21.4061
12.9532 | -34.5487 | -21.4147
12.9593 | -34.5491 | -21.4235
12.9654 | -34.5495 | -21.4323
12.9716 | -34.5499 | -21.4409
12.9777 | -34.5504 | -21.449
12.9838 | -34.5508 | -21.4575
12.9899 | -34.5513 | -21.466
12.996 | -34.5517 | -21.4747
13.0021 | -34.5521 | -21.4833
13.0082 | -34.5526 | -21.4918
13.0143 | -34.553 | -21.4999
13.0204 | -34.5534 | -21.5083
13.0265 | -34.5539 | -21.5165
13.0326 | -34.5543 | -21.5253
13.0387 | -34.5548 | -21.5339
13.0448 | -34.5552 | -21.5424
13.0509 | -34.5557 | -21.551
13.057 | -34.5562 | -21.5592
13.0631 | -34.5567 | -21.5675
13.0692 | -34.5572 | -21.5758
13.0753 | -34.5576 | -21.5842
13.0814 | -34.5581 | -21.5925
13.0875 | -34.5585 | -21.6006
13.0936 | -34.559 | -21.609
13.0997 | -34.5594 | -21.617
13.1058 | -34.5598 | -21.6249
13.112 | -34.5602 | -21.6334
13.1181 | -34.5607 | -21.6419
13.1242 | -34.5611 | -21.6504
13.1303 | -34.5616 | -21.6592
13.1364 | -34.562 | -21.6676
13.1425 | -34.5624 | -21.6759
13.1486 | -34.5629 | -21.6843
13.1547 | -34.5634 | -21.6924
13.1608 | -34.5638 | -21.7007
13.1669 | -34.5643 | -21.7092
13.173 | -34.5647 | -21.7176
13.1791 | -34.5652 | -21.7261
13.1852 | -34.5656 | -21.7347
13.1913 | -34.5661 | -21.7431
13.1974 | -34.5665 | -21.7516
13.2035 | -34.5669 | -21.76
13.2096 | -34.5674 | -21.7683
13.2157 | -34.5679 | -21.7767
13.2218 | -34.5683 | -21.7852
13.2279 | -34.5688 | -21.7935
13.234 | -34.5693 | -21.8019
13.2401 | -34.5697 | -21.8102
13.2462 | -34.5702 | -21.8185
13.2524 | -34.5706 | -21.8268
13.2585 | -34.5711 | -21.835
13.2646 | -34.5715 | -21.8435
13.2707 | -34.572 | -21.852
13.2768 | -34.5724 | -21.8608
13.2829 | -34.5728 | -21.8694
13.289 | -34.5733 | -21.8777
13.2951 | -34.5738 | -21.8857
13.3012 | -34.5742 | -21.8941
13.3073 | -34.5746 | -21.9027
13.3134 | -34.575 | -21.911
13.3195 | -34.5755 | -21.9192
13.3256 | -34.576 | -21.9277
13.3317 | -34.5764 | -21.9358
13.3378 | -34.5769 | -21.9443
13.3439 | -34.5773 | -21.9526
13.35 | -34.5778 | -21.9609
13.3561 | -34.5782 | -21.9694
13.3622 | -34.5787 | -21.9776
13.3683 | -34.5791 | -21.9857
13.3744 | -34.5796 | -21.994
13.3805 | -34.58 | -22.0024
13.3866 | -34.5805 | -22.0106
13.3927 | -34.5809 | -22.0189
13.3989 | -34.5813 | -22.0267
13.405 | -34.5818 | -22.0349
13.4111 | -34.5823 | -22.0431
13.4172 | -34.5827 | -22.0512
13.4233 | -34.5832 | -22.0596
13.4294 | -34.5836 | -22.0677
13.4355 | -34.5841 | -22.0767
13.4416 | -34.5845 | -22.085
13.4477 | -34.5849 | -22.0932
13.4538 | -34.5854 | -22.1016
13.4599 | -34.5859 | -22.1101
13.466 | -34.5863 | -22.1188
13.4721 | -34.5868 | -22.1272
13.4782 | -34.5872 | -22.1353
13.4843 | -34.5878 | -22.1438
13.4904 | -34.5882 | -22.1515
13.4965 | -34.5886 | -22.1601
13.5026 | -34.589 | -22.1683
13.5087 | -34.5895 | -22.1764
13.5148 | -34.5899 | -22.1848
13.5209 | -34.5904 | -22.1929
13.527 | -34.5908 | -22.2016
13.5331 | -34.5913 | -22.2098
13.5393 | -34.5917 | -22.218
13.5454 | -34.5921 | -22.2263
13.5515 | -34.5926 | -22.2347
13.5576 | -34.5931 | -22.243
13.5637 | -34.5935 | -22.2512
13.5698 | -34.594 | -22.2597
13.5759 | -34.5944 | -22.2676
13.582 | -34.5948 | -22.2761
13.5881 | -34.5953 | -22.2843
13.5942 | -34.5957 | -22.2927
13.6003 | -34.5962 | -22.3009
13.6064 | -34.5966 | -22.3094
13.6125 | -34.5971 | -22.3179
13.6186 | -34.5976 | -22.3261
13.6247 | -34.5981 | -22.3344
13.6308 | -34.5985 | -22.3425
13.6369 | -34.5989 | -22.3507
13.643 | -34.5994 | -22.359
13.6491 | -34.5998 | -22.3673
13.6552 | -34.6003 | -22.3753
13.6613 | -34.6008 | -22.3838
13.6674 | -34.6012 | -22.3922
13.6735 | -34.6017 | -22.4005
13.6796 | -34.6021 | -22.4086
13.6858 | -34.6025 | -22.4169
13.6919 | -34.603 | -22.4253
13.698 | -34.6034 | -22.4336
13.7041 | -34.6038 | -22.4421
13.7102 | -34.6043 | -22.4505
13.7163 | -34.6048 | -22.4589
13.7224 | -34.6053 | -22.4671
13.7285 | -34.6057 | -22.4755
13.7346 | -34.6062 | -22.4838
13.7407 | -34.6066 | -22.492
13.7468 | -34.6071 | -22.5003
13.7529 | -34.6075 | -22.5084
13.759 | -34.6079 | -22.5168
13.7651 | -34.6084 | -22.5252
13.7712 | -34.6088 | -22.5336
13.7773 | -34.6093 | -22.542
13.7834 | -34.6098 | -22.5502
13.7895 | -34.6103 | -22.5583
13.7956 | -34.6107 | -22.5667
13.8017 | -34.6112 | -22.5748
13.8078 | -34.6117 | -22.5831
13.8139 | -34.6122 | -22.5912
13.82 | -34.6126 | -22.5992
13.8262 | -34.6131 | -22.6078
13.8323 | -34.6136 | -22.6162
13.8384 | -34.614 | -22.6247
13.8445 | -34.6145 | -22.6331
13.8506 | -34.6149 | -22.6414
13.8567 | -34.6153 | -22.65
13.8628 | -34.6158 | -22.6584
13.8689 | -34.6162 | -22.6667
13.875 | -34.6167 | -22.6749
13.8811 | -34.6172 | -22.6831
13.8872 | -34.6177 | -22.6913
13.8933 | -34.6181 | -22.6993
13.8994 | -34.6186 | -22.7075
13.9055 | -34.6191 | -22.7157
13.9116 | -34.6195 | -22.7238
13.9177 | -34.62 | -22.7317
13.9238 | -34.6205 | -22.74
13.9299 | -34.621 | -22.7483
13.936 | -34.6214 | -22.7563
13.9421 | -34.6219 | -22.7646
13.9482 | -34.6224 | -22.773
13.9543 | -34.6229 | -22.781
13.9604 | -34.6233 | -22.7894
13.9665 | -34.6238 | -22.7977
13.9727 | -34.6242 | -22.8059
13.9788 | -34.6247 | -22.8143
13.9849 | -34.6251 | -22.8226
13.991 | -34.6255 | -22.8312
13.9971 | -34.626 | -22.8397
14.0032 | -34.6264 | -22.8481
14.0093 | -34.6268 | -22.8562
14.0154 | -34.6272 | -22.8647
14.0215 | -34.6277 | -22.8726
14.0276 | -34.6281 | -22.8807
14.0337 | -34.6287 | -22.8891
14.0398 | -34.6291 | -22.8971
14.0459 | -34.6295 | -22.9055
14.052 | -34.6299 | -22.9137
14.0581 | -34.6304 | -22.9216
14.0642 | -34.6309 | -22.93
14.0703 | -34.6313 | -22.9383
14.0764 | -34.6318 | -22.9469
14.0825 | -34.6323 | -22.9551
14.0886 | -34.6327 | -22.9629
14.0947 | -34.6331 | -22.9709
14.1008 | -34.6336 | -22.9795
14.1069 | -34.634 | -22.9876
14.1131 | -34.6345 | -22.9961
14.1192 | -34.6349 | -23.004
14.1253 | -34.6353 | -23.0121
14.1314 | -34.6358 | -23.0202
14.1375 | -34.6363 | -23.0287
14.1436 | -34.6368 | -23.0369
14.1497 | -34.6373 | -23.0449
14.1558 | -34.6377 | -23.0533
14.1619 | -34.6382 | -23.0614
14.168 | -34.6386 | -23.0696
14.1741 | -34.6392 | -23.0777
14.1802 | -34.6396 | -23.0858
14.1863 | -34.64 | -23.094
14.1924 | -34.6405 | -23.1023
14.1985 | -34.641 | -23.1102
14.2046 | -34.6414 | -23.1183
14.2107 | -34.6419 | -23.1261
14.2168 | -34.6424 | -23.1347
14.2229 | -34.6428 | -23.1428
14.229 | -34.6433 | -23.151
14.2351 | -34.6437 | -23.159
14.2412 | -34.6442 | -23.1669
14.2473 | -34.6447 | -23.1752
14.2534 | -34.6452 | -23.1832
14.2596 | -34.6456 | -23.1914
14.2657 | -34.6461 | -23.1997
14.2718 | -34.6466 | -23.208
14.2779 | -34.647 | -23.216
14.284 | -34.6475 | -23.2241
14.2901 | -34.648 | -23.2321
14.2962 | -34.6485 | -23.24
14.3023 | -34.6489 | -23.2483
14.3084 | -34.6494 | -23.2567
14.3145 | -34.6499 | -23.265
14.3206 | -34.6503 | -23.2733
14.3267 | -34.6508 | -23.2815
14.3328 | -34.6512 | -23.2895
14.3389 | -34.6517 | -23.2974
14.345 | -34.6521 | -23.3058
14.3511 | -34.6526 | -23.3137
14.3572 | -34.653 | -23.322
14.3633 | -34.6535 | -23.3301
14.3694 | -34.6539 | -23.3382
14.3755 | -34.6543 | -23.3461
14.3816 | -34.6548 | -23.3542
14.3877 | -34.6553 | -23.3622
14.3938 | -34.6557 | -23.3706
14.4 | -34.6562 | -23.3785
14.4061 | -34.6566 | -23.3867
14.4122 | -34.657 | -23.395
14.4183 | -34.6574 | -23.4032
14.4244 | -34.6579 | -23.4112
14.4305 | -34.6584 | -23.4196
14.4366 | -34.6588 | -23.4278
14.4427 | -34.6593 | -23.4361
14.4488 | -34.6597 | -23.4439
14.4549 | -34.6602 | -23.452
14.461 | -34.6607 | -23.4603
14.4671 | -34.6611 | -23.4685
14.4732 | -34.6616 | -23.4763
14.4793 | -34.6621 | -23.4852
14.4854 | -34.6626 | -23.493
14.4915 | -34.663 | -23.501
14.4976 | -34.6636 | -23.5093
14.5037 | -34.6641 | -23.5175
14.5098 | -34.6645 | -23.5253
14.5159 | -34.665 | -23.5336
14.522 | -34.6654 | -23.5418
14.5281 | -34.6659 | -23.5497
14.5342 | -34.6663 | -23.558
14.5403 | -34.6668 | -23.5661
14.5465 | -34.6673 | -23.5745
14.5526 | -34.6678 | -23.583
14.5587 | -34.6682 | -23.5911
14.5648 | -34.6687 | -23.599
14.5709 | -34.6692 | -23.6073
14.577 | -34.6696 | -23.6153
14.5831 | -34.6701 | -23.6236
14.5892 | -34.6706 | -23.6314
14.5953 | -34.6711 | -23.6392
14.6014 | -34.6715 | -23.6474
14.6075 | -34.672 | -23.6553
14.6136 | -34.6725 | -23.6632
14.6197 | -34.6729 | -23.6714
14.6258 | -34.6734 | -23.6795
14.6319 | -34.6739 | -23.6874
14.638 | -34.6744 | -23.6956
14.6441 | -34.6748 | -23.7033
14.6502 | -34.6753 | -23.7111
14.6563 | -34.6757 | -23.7192
14.6624 | -34.6762 | -23.7273
14.6685 | -34.6767 | -23.735
14.6746 | -34.6772 | -23.7431
14.6807 | -34.6776 | -23.7513
14.6869 | -34.678 | -23.7592
14.693 | -34.6785 | -23.7671
14.6991 | -34.679 | -23.7751
14.7052 | -34.6795 | -23.7831
14.7113 | -34.6799 | -23.7914
14.7174 | -34.6804 | -23.7995
14.7235 | -34.6808 | -23.8074
14.7296 | -34.6813 | -23.8153
14.7357 | -34.6818 | -23.8233
14.7418 | -34.6822 | -23.8311
14.7479 | -34.6827 | -23.8391
14.754 | -34.6832 | -23.8471
14.7601 | -34.6837 | -23.8553
14.7662 | -34.6841 | -23.8634
14.7723 | -34.6846 | -23.8714
14.7784 | -34.685 | -23.8795
14.7845 | -34.6855 | -23.8877
14.7906 | -34.686 | -23.8958
14.7967 | -34.6864 | -23.9038
14.8028 | -34.6869 | -23.9116
14.8089 | -34.6874 | -23.9196
14.815 | -34.6878 | -23.9274
14.8211 | -34.6883 | -23.9353
14.8272 | -34.6887 | -23.9436
14.8334 | -34.6892 | -23.9515
14.8395 | -34.6897 | -23.9595
14.8456 | -34.6902 | -23.9675
14.8517 | -34.6906 | -23.9756
14.8578 | -34.6911 | -23.9836
14.8639 | -34.6916 | -23.9919
14.87 | -34.6921 | -24.0001
14.8761 | -34.6926 | -24.0079
14.8822 | -34.693 | -24.0161
14.8883 | -34.6935 | -24.024
14.8944 | -34.694 | -24.032
14.9005 | -34.6945 | -24.0398
14.9066 | -34.6949 | -24.0479
14.9127 | -34.6954 | -24.0558
14.9188 | -34.6959 | -24.0641
14.9249 | -34.6963 | -24.0721
14.931 | -34.6968 | -24.0799
14.9371 | -34.6973 | -24.0879
14.9432 | -34.6978 | -24.096
14.9493 | -34.6982 | -24.104
14.9554 | -34.6987 | -24.1119
14.9615 | -34.6992 | -24.12
14.9676 | -34.6996 | -24.1282
14.9738 | -34.7001 | -24.1359
14.9799 | -34.7005 | -24.1444
14.986 | -34.701 | -24.1521
14.9921 | -34.7016 | -24.16
14.9982 | -34.702 | -24.1684
15.0043 | -34.7024 | -24.1763
15.0104 | -34.7029 | -24.1843
15.0165 | -34.7034 | -24.1924
15.0226 | -34.7038 | -24.2
15.0287 | -34.7043 | -24.2084
15.0348 | -34.7048 | -24.2166
15.0409 | -34.7052 | -24.2248
15.047 | -34.7057 | -24.233
15.0531 | -34.7063 | -24.2411
15.0592 | -34.7068 | -24.2493
15.0653 | -34.7073 | -24.2573
15.0714 | -34.7077 | -24.2651
15.0775 | -34.7083 | -24.2729
15.0836 | -34.7088 | -24.2809
15.0897 | -34.7093 | -24.2888
15.0958 | -34.7097 | -24.2967
15.1019 | -34.7102 | -24.3049
15.108 | -34.7106 | -24.3128
15.1141 | -34.7111 | -24.3207
15.1203 | -34.7116 | -24.3287
15.1264 | -34.7121 | -24.3369
15.1325 | -34.7126 | -24.3447
15.1386 | -34.7131 | -24.3528
15.1447 | -34.7136 | -24.3608
15.1508 | -34.7141 | -24.3691
15.1569 | -34.7146 | -24.3773
15.163 | -34.7151 | -24.3848
15.1691 | -34.7156 | -24.3927
15.1752 | -34.7161 | -24.4006
15.1813 | -34.7166 | -24.4085
15.1874 | -34.7171 | -24.4166
15.1935 | -34.7176 | -24.4243
15.1996 | -34.7181 | -24.4324
15.2057 | -34.7186 | -24.4402
15.2118 | -34.719 | -24.4482
15.2179 | -34.7196 | -24.4559
15.224 | -34.72 | -24.4641
15.2301 | -34.7205 | -24.4718
15.2362 | -34.721 | -24.4798
15.2423 | -34.7215 | -24.4879
15.2484 | -34.722 | -24.4956
15.2545 | -34.7225 | -24.5034
15.2607 | -34.723 | -24.5115
15.2668 | -34.7235 | -24.5193
15.2729 | -34.724 | -24.5273
15.279 | -34.7245 | -24.5353
15.2851 | -34.725 | -24.5434
15.2912 | -34.7254 | -24.5512
15.2973 | -34.7259 | -24.5593
15.3034 | -34.7264 | -24.567
15.3095 | -34.7269 | -24.5747
15.3156 | -34.7273 | -24.5827
15.3217 | -34.7278 | -24.5901
15.3278 | -34.7284 | -24.5982
15.3339 | -34.7289 | -24.6059
15.34 | -34.7294 | -24.6137
15.3461 | -34.7298 | -24.6217
15.3522 | -34.7303 | -24.6296
15.3583 | -34.7308 | -24.6374
15.3644 | -34.7313 | -24.6455
15.3705 | -34.7318 | -24.6532
15.3766 | -34.7323 | -24.6613
15.3827 | -34.7327 | -24.669
15.3888 | -34.7332 | -24.6766
15.3949 | -34.7337 | -24.6844
15.401 | -34.7342 | -24.6924
15.4072 | -34.7347 | -24.7
15.4133 | -34.7352 | -24.708
15.4194 | -34.7357 | -24.7161
15.4255 | -34.7361 | -24.724
15.4316 | -34.7366 | -24.7319
15.4377 | -34.7371 | -24.7396
15.4438 | -34.7377 | -24.7473
15.4499 | -34.7381 | -24.7549
15.456 | -34.7386 | -24.7628
15.4621 | -34.7391 | -24.7708
15.4682 | -34.7396 | -24.7789
15.4743 | -34.7402 | -24.7866
15.4804 | -34.7406 | -24.7946
15.4865 | -34.7411 | -24.8024
15.4926 | -34.7416 | -24.8104
15.4987 | -34.7421 | -24.8183
15.5048 | -34.7426 | -24.8262
15.5109 | -34.7431 | -24.8337
15.517 | -34.7436 | -24.8418
15.5231 | -34.7441 | -24.8497
15.5292 | -34.7446 | -24.8573
15.5353 | -34.7451 | -24.8652
15.5414 | -34.7456 | -24.8733
15.5476 | -34.7461 | -24.8813
15.5537 | -34.7466 | -24.8891
15.5598 | -34.7472 | -24.8966
15.5659 | -34.7477 | -24.9047
15.572 | -34.7482 | -24.9122
15.5781 | -34.7486 | -24.9202
15.5842 | -34.7492 | -24.9278
15.5903 | -34.7497 | -24.9354
15.5964 | -34.7502 | -24.9432
15.6025 | -34.7507 | -24.9507
15.6086 | -34.7512 | -24.9586
15.6147 | -34.7517 | -24.9665
15.6208 | -34.7522 | -24.9742
15.6269 | -34.7528 | -24.9819
15.633 | -34.7533 | -24.9898
15.6391 | -34.7537 | -24.9977
15.6452 | -34.7543 | -25.0055
15.6513 | -34.7548 | -25.0137
15.6574 | -34.7552 | -25.0213
15.6635 | -34.7557 | -25.0292
15.6696 | -34.7562 | -25.0368
15.6757 | -34.7567 | -25.0448
15.6818 | -34.7572 | -25.0529
15.688 | -34.7578 | -25.0607
15.6941 | -34.7582 | -25.0687
15.7002 | -34.7587 | -25.0764
15.7063 | -34.7592 | -25.0845
15.7124 | -34.7597 | -25.0925
15.7185 | -34.7602 | -25.1003
15.7246 | -34.7607 | -25.1079
15.7307 | -34.7612 | -25.1159
15.7368 | -34.7618 | -25.124
15.7429 | -34.7623 | -25.1318
15.749 | -34.7628 | -25.1397
15.7551 | -34.7632 | -25.1476
15.7612 | -34.7637 | -25.155
15.7673 | -34.7642 | -25.1627
15.7734 | -34.7647 | -25.1703
15.7795 | -34.7652 | -25.1786
15.7856 | -34.7657 | -25.1861
15.7917 | -34.7662 | -25.1938
15.7978 | -34.7668 | -25.2016
15.8039 | -34.7672 | -25.2094
15.81 | -34.7677 | -25.2171
15.8161 | -34.7682 | -25.225
15.8222 | -34.7687 | -25.2329
15.8283 | -34.7693 | -25.2407
15.8345 | -34.7698 | -25.2485
15.8406 | -34.7703 | -25.2563
15.8467 | -34.7708 | -25.2642
15.8528 | -34.7714 | -25.272
15.8589 | -34.7718 | -25.2797
15.865 | -34.7723 | -25.2875
15.8711 | -34.7729 | -25.2955
15.8772 | -34.7734 | -25.3034
15.8833 | -34.774 | -25.3116
15.8894 | -34.7744 | -25.3194
15.8955 | -34.7749 | -25.3272
15.9016 | -34.7755 | -25.335
15.9077 | -34.776 | -25.3428
15.9138 | -34.7764 | -25.3505
15.9199 | -34.777 | -25.3584
15.926 | -34.7775 | -25.3665
15.9321 | -34.778 | -25.3743
15.9382 | -34.7786 | -25.3819
15.9443 | -34.7791 | -25.3899
15.9504 | -34.7796 | -25.3976
15.9565 | -34.7802 | -25.4053
15.9626 | -34.7807 | -25.4131
15.9687 | -34.7812 | -25.4207
15.9749 | -34.7818 | -25.4286
15.981 | -34.7823 | -25.4362
15.9871 | -34.7828 | -25.4439
15.9932 | -34.7832 | -25.4518
15.9993 | -34.7837 | -25.4593
16.0054 | -34.7843 | -25.4669
16.0115 | -34.7848 | -25.4746
16.0176 | -34.7853 | -25.4826
16.0237 | -34.7859 | -25.4907
16.0298 | -34.7864 | -25.498
16.0359 | -34.7869 | -25.5058
16.042 | -34.7874 | -25.5136
16.0481 | -34.7879 | -25.5214
16.0542 | -34.7884 | -25.5293
16.0603 | -34.789 | -25.537
16.0664 | -34.7895 | -25.5449
16.0725 | -34.79 | -25.5525
16.0786 | -34.7906 | -25.5605
16.0847 | -34.791 | -25.5681
16.0908 | -34.7915 | -25.5758
16.0969 | -34.7921 | -25.5834
16.103 | -34.7925 | -25.5914
16.1091 | -34.7931 | -25.5992
16.1152 | -34.7936 | -25.6069
16.1214 | -34.7942 | -25.6147
16.1275 | -34.7946 | -25.6223
16.1336 | -34.7952 | -25.63
16.1397 | -34.7957 | -25.6379
16.1458 | -34.7962 | -25.6455
16.1519 | -34.7967 | -25.6534
16.158 | -34.7972 | -25.6613
16.1641 | -34.7977 | -25.669
16.1702 | -34.7982 | -25.6768
16.1763 | -34.7988 | -25.6846
16.1824 | -34.7993 | -25.6923
16.1885 | -34.7998 | -25.6998
16.1946 | -34.8003 | -25.7074
16.2007 | -34.8008 | -25.7157
16.2068 | -34.8013 | -25.7233
16.2129 | -34.8018 | -25.7311
16.219 | -34.8023 | -25.739
16.2251 | -34.8028 | -25.747
16.2312 | -34.8033 | -25.7548
16.2373 | -34.8038 | -25.7625
16.2434 | -34.8043 | -25.7701
16.2495 | -34.8049 | -25.7777
16.2556 | -34.8054 | -25.7855
16.2618 | -34.8059 | -25.7929
16.2679 | -34.8065 | -25.8007
16.274 | -34.807 | -25.8085
16.2801 | -34.8075 | -25.8161
16.2862 | -34.8081 | -25.8237
16.2923 | -34.8085 | -25.8317
16.2984 | -34.809 | -25.8396
16.3045 | -34.8096 | -25.8473
16.3106 | -34.8101 | -25.8547
16.3167 | -34.8106 | -25.8627
16.3228 | -34.8112 | -25.8705
16.3289 | -34.8116 | -25.8784
16.335 | -34.8121 | -25.8863
16.3411 | -34.8126 | -25.8938
16.3472 | -34.8132 | -25.9016
16.3533 | -34.8137 | -25.9094
16.3594 | -34.8142 | -25.917
16.3655 | -34.8147 | -25.925
16.3716 | -34.8152 | -25.9326
16.3777 | -34.8157 | -25.9404
16.3838 | -34.8162 | -25.9481
16.3899 | -34.8167 | -25.956
16.396 | -34.8172 | -25.9639
16.4021 | -34.8178 | -25.9716
16.4083 | -34.8183 | -25.9793
16.4144 | -34.8189 | -25.987
16.4205 | -34.8194 | -25.9949
16.4266 | -34.8199 | -26.0026
16.4327 | -34.8204 | -26.0101
16.4388 | -34.8209 | -26.0179
16.4449 | -34.8215 | -26.0256
16.451 | -34.822 | -26.0335
16.4571 | -34.8224 | -26.0415
16.4632 | -34.823 | -26.0489
16.4693 | -34.8234 | -26.0569
16.4754 | -34.824 | -26.0647
16.4815 | -34.8245 | -26.0725
16.4876 | -34.825 | -26.0802
16.4937 | -34.8255 | -26.088
16.4998 | -34.8261 | -26.0957
16.5059 | -34.8266 | -26.1035
16.512 | -34.8271 | -26.1115
16.5181 | -34.8276 | -26.1195
16.5242 | -34.8281 | -26.1273
16.5303 | -34.8286 | -26.135
16.5364 | -34.8291 | -26.1426
16.5425 | -34.8297 | -26.1505
16.5487 | -34.8302 | -26.158
16.5548 | -34.8306 | -26.1658
16.5609 | -34.8312 | -26.1737
16.567 | -34.8317 | -26.1813
16.5731 | -34.8322 | -26.189
16.5792 | -34.8328 | -26.1968
16.5853 | -34.8333 | -26.2044
16.5914 | -34.8338 | -26.2121
16.5975 | -34.8343 | -26.2199
16.6036 | -34.8349 | -26.2276
16.6097 | -34.8354 | -26.2352
16.6158 | -34.8358 | -26.2431
16.6219 | -34.8364 | -26.2508
16.628 | -34.8369 | -26.2585
16.6341 | -34.8374 | -26.266
16.6402 | -34.8379 | -26.2731
16.6463 | -34.8384 | -26.2809
16.6524 | -34.8389 | -26.2884
16.6585 | -34.8394 | -26.2961
16.6646 | -34.8399 | -26.3038
16.6707 | -34.8405 | -26.3112
16.6768 | -34.841 | -26.3188
16.6829 | -34.8415 | -26.3265
16.689 | -34.842 | -26.3342
16.6952 | -34.8425 | -26.342
16.7013 | -34.843 | -26.3498
16.7074 | -34.8436 | -26.3574
16.7135 | -34.8441 | -26.3652
16.7196 | -34.8446 | -26.3729
16.7257 | -34.8451 | -26.3806
16.7318 | -34.8456 | -26.3884
16.7379 | -34.8461 | -26.396
16.744 | -34.8467 | -26.4036
16.7501 | -34.8471 | -26.4113
16.7562 | -34.8477 | -26.419
16.7623 | -34.8482 | -26.4268
16.7684 | -34.8487 | -26.4345
16.7745 | -34.8492 | -26.4419
16.7806 | -34.8497 | -26.4497
16.7867 | -34.8503 | -26.4572
16.7928 | -34.8508 | -26.4649
16.7989 | -34.8512 | -26.4725
16.805 | -34.8517 | -26.4803
16.8111 | -34.8522 | -26.488
16.8172 | -34.8528 | -26.4959
16.8233 | -34.8533 | -26.5037
16.8294 | -34.8539 | -26.5115
16.8356 | -34.8544 | -26.5194
16.8417 | -34.8549 | -26.5269
16.8478 | -34.8554 | -26.5344
16.8539 | -34.8559 | -26.5419
16.86 | -34.8564 | -26.5496
16.8661 | -34.8569 | -26.5574
16.8722 | -34.8575 | -26.5652
16.8783 | -34.858 | -26.5729
16.8844 | -34.8585 | -26.5803
16.8905 | -34.8591 | -26.588
16.8966 | -34.8596 | -26.5959
16.9027 | -34.8601 | -26.6031
16.9088 | -34.8606 | -26.6112
16.9149 | -34.8612 | -26.619
16.921 | -34.8617 | -26.6265
16.9271 | -34.8622 | -26.6344
16.9332 | -34.8627 | -26.6421
16.9393 | -34.8632 | -26.65
16.9454 | -34.8637 | -26.6574
16.9515 | -34.8642 | -26.6654
16.9576 | -34.8647 | -26.6728
16.9637 | -34.8653 | -26.6803
16.9698 | -34.8658 | -26.6876
16.9759 | -34.8662 | -26.6952
16.9821 | -34.8667 | -26.7028
16.9882 | -34.8672 | -26.7106
16.9943 | -34.8678 | -26.7186
17.0004 | -34.8683 | -26.7261
17.0065 | -34.8688 | -26.7339
17.0126 | -34.8693 | -26.7416
17.0187 | -34.8698 | -26.7494
17.0248 | -34.8703 | -26.7568
17.0309 | -34.8708 | -26.7645
17.037 | -34.8714 | -26.7721
17.0431 | -34.8719 | -26.7794
17.0492 | -34.8724 | -26.7871
17.0553 | -34.8729 | -26.795
17.0614 | -34.8734 | -26.8022
17.0675 | -34.8739 | -26.8103
17.0736 | -34.8745 | -26.8178
17.0797 | -34.875 | -26.8251
17.0858 | -34.8755 | -26.8328
17.0919 | -34.876 | -26.8407
17.098 | -34.8765 | -26.8481
17.1041 | -34.877 | -26.8555
17.1102 | -34.8776 | -26.8633
17.1163 | -34.8781 | -26.8709
17.1225 | -34.8786 | -26.8783
17.1286 | -34.8791 | -26.8855
17.1347 | -34.8796 | -26.893
17.1408 | -34.8802 | -26.9005
17.1469 | -34.8807 | -26.9081
17.153 | -34.8812 | -26.9157
17.1591 | -34.8817 | -26.9232
17.1652 | -34.8823 | -26.9305
17.1713 | -34.8828 | -26.9378
17.1774 | -34.8833 | -26.9456
17.1835 | -34.8839 | -26.9533
17.1896 | -34.8844 | -26.9607
17.1957 | -34.8849 | -26.9684
17.2018 | -34.8854 | -26.9758
17.2079 | -34.8859 | -26.9832
17.214 | -34.8863 | -26.9907
17.2201 | -34.8868 | -26.9985
17.2262 | -34.8873 | -27.0061
17.2323 | -34.8878 | -27.0136
17.2384 | -34.8883 | -27.0211
17.2445 | -34.8888 | -27.0285
17.2506 | -34.8893 | -27.0361
17.2567 | -34.8898 | -27.0439
17.2628 | -34.8903 | -27.0518
17.269 | -34.8908 | -27.0594
17.2751 | -34.8913 | -27.0669
17.2812 | -34.8919 | -27.0745
17.2873 | -34.8924 | -27.0819
17.2934 | -34.8929 | -27.0894
17.2995 | -34.8933 | -27.0967
17.3056 | -34.8938 | -27.1041
17.3117 | -34.8943 | -27.1116
17.3178 | -34.8948 | -27.119
17.3239 | -34.8953 | -27.1264
17.33 | -34.8958 | -27.1337
17.3361 | -34.8963 | -27.1411
17.3422 | -34.8968 | -27.1486
17.3483 | -34.8973 | -27.1562
17.3544 | -34.8978 | -27.1637
17.3605 | -34.8983 | -27.1712
17.3666 | -34.8989 | -27.179
17.3727 | -34.8994 | -27.1863
17.3788 | -34.8998 | -27.1936
17.3849 | -34.9003 | -27.2011
17.391 | -34.9009 | -27.2087
17.3971 | -34.9014 | -27.2163
17.4032 | -34.9019 | -27.2243
17.4094 | -34.9024 | -27.2315
17.4155 | -34.9029 | -27.2389
17.4216 | -34.9034 | -27.2463
17.4277 | -34.9039 | -27.254
17.4338 | -34.9044 | -27.2615
17.4399 | -34.9049 | -27.2691
17.446 | -34.9054 | -27.2762
17.4521 | -34.9059 | -27.2835
17.4582 | -34.9064 | -27.2907
17.4643 | -34.9069 | -27.2985
17.4704 | -34.9073 | -27.306
17.4765 | -34.9078 | -27.3133
17.4826 | -34.9084 | -27.3205
17.4887 | -34.9089 | -27.328
17.4948 | -34.9094 | -27.3352
17.5009 | -34.9098 | -27.3424
17.507 | -34.9104 | -27.3501
17.5131 | -34.9109 | -27.3575
17.5192 | -34.9113 | -27.3649
17.5253 | -34.9118 | -27.3725
17.5314 | -34.9122 | -27.38
17.5375 | -34.9127 | -27.3875
17.5436 | -34.9132 | -27.3952
17.5497 | -34.9137 | -27.4027
17.5559 | -34.9142 | -27.4099
17.562 | -34.9147 | -27.4173
17.5681 | -34.9152 | -27.4247
17.5742 | -34.9157 | -27.432
17.5803 | -34.9162 | -27.4394
17.5864 | -34.9168 | -27.4469
17.5925 | -34.9174 | -27.4547
17.5986 | -34.9179 | -27.4617
17.6047 | -34.9184 | -27.4695
17.6108 | -34.9189 | -27.4769
17.6169 | -34.9194 | -27.484
17.623 | -34.9199 | -27.4911
17.6291 | -34.9205 | -27.4985
17.6352 | -34.921 | -27.5059
17.6413 | -34.9215 | -27.5137
17.6474 | -34.9219 | -27.5214
17.6535 | -34.9224 | -27.5287
17.6596 | -34.9229 | -27.536
17.6657 | -34.9234 | -27.5433
17.6718 | -34.9239 | -27.551
17.6779 | -34.9244 | -27.5583
17.684 | -34.9249 | -27.5657
17.6901 | -34.9254 | -27.5731
17.6963 | -34.9259 | -27.5805
17.7024 | -34.9264 | -27.5876
17.7085 | -34.9269 | -27.5949
17.7146 | -34.9275 | -27.6023
17.7207 | -34.928 | -27.6099
17.7268 | -34.9285 | -27.6175
17.7329 | -34.929 | -27.6251
17.739 | -34.9294 | -27.6327
17.7451 | -34.93 | -27.6399
17.7512 | -34.9305 | -27.6471
17.7573 | -34.931 | -27.6546
17.7634 | -34.9315 | -27.6618
17.7695 | -34.932 | -27.6692
17.7756 | -34.9325 | -27.6765
17.7817 | -34.933 | -27.6835
17.7878 | -34.9335 | -27.6907
17.7939 | -34.9339 | -27.6979
17.8 | -34.9345 | -27.7052
17.8061 | -34.9349 | -27.7126
17.8122 | -34.9354 | -27.7201
17.8183 | -34.936 | -27.7277
17.8244 | -34.9364 | -27.7349
17.8305 | -34.9369 | -27.7423
17.8366 | -34.9374 | -27.7491
17.8428 | -34.9379 | -27.7565
17.8489 | -34.9383 | -27.7637
17.855 | -34.9389 | -27.7708
17.8611 | -34.9394 | -27.7779
17.8672 | -34.9399 | -27.7853
17.8733 | -34.9404 | -27.7927
17.8794 | -34.9409 | -27.8002
17.8855 | -34.9414 | -27.8076
17.8916 | -34.9419 | -27.8147
17.8977 | -34.9424 | -27.822
17.9038 | -34.9429 | -27.8289
17.9099 | -34.9434 | -27.8361
17.916 | -34.9439 | -27.8434
17.9221 | -34.9444 | -27.8509
17.9282 | -34.9449 | -27.858
17.9343 | -34.9454 | -27.8654
17.9404 | -34.9459 | -27.8725
17.9465 | -34.9464 | -27.8799
17.9526 | -34.9469 | -27.8873
17.9587 | -34.9474 | -27.8946
17.9648 | -34.9479 | -27.9017
17.9709 | -34.9484 | -27.9089
17.977 | -34.9489 | -27.916
17.9832 | -34.9494 | -27.9235
17.9893 | -34.9499 | -27.931
17.9954 | -34.9504 | -27.9385
18.0015 | -34.9509 | -27.946
18.0076 | -34.9513 | -27.9526
18.0137 | -34.9518 | -27.9602
18.0198 | -34.9523 | -27.9675
18.0259 | -34.9528 | -27.975
18.032 | -34.9533 | -27.9822
18.0381 | -34.9538 | -27.9893
18.0442 | -34.9543 | -27.9965
18.0503 | -34.9547 | -28.0041
18.0564 | -34.9552 | -28.0113
18.0625 | -34.9557 | -28.0188
18.0686 | -34.9562 | -28.026
18.0747 | -34.9567 | -28.0332
18.0808 | -34.9572 | -28.0401
18.0869 | -34.9577 | -28.0474
18.093 | -34.9583 | -28.0546
18.0991 | -34.9587 | -28.0621
18.1052 | -34.9592 | -28.0697
18.1113 | -34.9597 | -28.0766
18.1174 | -34.9602 | -28.0835
18.1236 | -34.9607 | -28.091
18.1297 | -34.9612 | -28.0981
18.1358 | -34.9617 | -28.1055
18.1419 | -34.9623 | -28.113
18.148 | -34.9628 | -28.1202
18.1541 | -34.9634 | -28.1271
18.1602 | -34.9639 | -28.1345
18.1663 | -34.9644 | -28.1416
18.1724 | -34.9649 | -28.1491
18.1785 | -34.9654 | -28.1565
18.1846 | -34.9659 | -28.1634
18.1907 | -34.9664 | -28.1709
18.1968 | -34.9669 | -28.1783
18.2029 | -34.9674 | -28.1851
18.209 | -34.9679 | -28.1926
18.2151 | -34.9684 | -28.2
18.2212 | -34.9689 | -28.2068
18.2273 | -34.9695 | -28.2139
18.2334 | -34.97 | -28.2214
18.2395 | -34.9704 | -28.2286
18.2456 | -34.9709 | -28.2359
18.2517 | -34.9714 | -28.2432
18.2578 | -34.9719 | -28.2501
18.2639 | -34.9724 | -28.2576
18.2701 | -34.9729 | -28.2649
18.2762 | -34.9734 | -28.2722
18.2823 | -34.974 | -28.2795
18.2884 | -34.9744 | -28.2869
18.2945 | -34.975 | -28.2941
18.3006 | -34.9755 | -28.3014
18.3067 | -34.976 | -28.3084
18.3128 | -34.9765 | -28.3157
18.3189 | -34.977 | -28.3233
18.325 | -34.9775 | -28.3297
18.3311 | -34.978 | -28.3369
18.3372 | -34.9785 | -28.344
18.3433 | -34.979 | -28.351
18.3494 | -34.9794 | -28.3579
18.3555 | -34.9799 | -28.3648
18.3616 | -34.9804 | -28.372
18.3677 | -34.981 | -28.3791
18.3738 | -34.9815 | -28.3861
18.3799 | -34.9819 | -28.393
18.386 | -34.9824 | -28.4
18.3921 | -34.9829 | -28.4072
18.3982 | -34.9834 | -28.4139
18.4043 | -34.9838 | -28.4209
18.4105 | -34.9843 | -28.428
18.4166 | -34.9848 | -28.4353
18.4227 | -34.9853 | -28.4426
18.4288 | -34.9858 | -28.4497
18.4349 | -34.9863 | -28.4567
18.441 | -34.9868 | -28.464
18.4471 | -34.9874 | -28.4712
18.4532 | -34.9879 | -28.4783
18.4593 | -34.9884 | -28.4854
18.4654 | -34.9888 | -28.4921
18.4715 | -34.9894 | -28.4995
18.4776 | -34.9899 | -28.5069
18.4837 | -34.9904 | -28.5141
18.4898 | -34.9909 | -28.5216
18.4959 | -34.9914 | -28.5287
18.502 | -34.992 | -28.5358
18.5081 | -34.9925 | -28.5431
18.5142 | -34.9929 | -28.5505
18.5203 | -34.9934 | -28.5573
18.5264 | -34.9939 | -28.5643
18.5325 | -34.9945 | -28.5715
18.5386 | -34.995 | -28.5789
18.5447 | -34.9955 | -28.5856
18.5508 | -34.996 | -28.5924
18.557 | -34.9965 | -28.5998
18.5631 | -34.997 | -28.607
18.5692 | -34.9975 | -28.6142
18.5753 | -34.998 | -28.6209
18.5814 | -34.9985 | -28.6282
18.5875 | -34.999 | -28.6356
18.5936 | -34.9995 | -28.6429
18.5997 | -35 | -28.6502
18.6058 | -35.0005 | -28.6575
18.6119 | -35.001 | -28.6645
18.618 | -35.0015 | -28.672
18.6241 | -35.002 | -28.6796
18.6302 | -35.0025 | -28.687
18.6363 | -35.003 | -28.6936
18.6424 | -35.0035 | -28.7005
18.6485 | -35.004 | -28.7077
18.6546 | -35.0044 | -28.715
18.6607 | -35.0049 | -28.722
18.6668 | -35.0055 | -28.7291
18.6729 | -35.0059 | -28.7362
18.679 | -35.0064 | -28.7432
18.6851 | -35.0069 | -28.7501
18.6912 | -35.0074 | -28.7572
18.6974 | -35.0079 | -28.7644
18.7035 | -35.0084 | -28.7715
18.7096 | -35.0089 | -28.7787
18.7157 | -35.0094 | -28.7857
18.7218 | -35.0099 | -28.7929
18.7279 | -35.0105 | -28.7997
18.734 | -35.011 | -28.8071
18.7401 | -35.0115 | -28.8138
18.7462 | -35.0119 | -28.8204
18.7523 | -35.0124 | -28.8274
18.7584 | -35.0129 | -28.8343
18.7645 | -35.0135 | -28.8414
18.7706 | -35.0139 | -28.8486
18.7767 | -35.0144 | -28.8558
18.7828 | -35.015 | -28.8625
18.7889 | -35.0154 | -28.8693
18.795 | -35.0159 | -28.8763
18.8011 | -35.0164 | -28.8833
18.8072 | -35.0169 | -28.8906
18.8133 | -35.0174 | -28.8974
18.8194 | -35.0177 | -28.9043
18.8255 | -35.0183 | -28.9112
18.8316 | -35.0188 | -28.9185
18.8377 | -35.0192 | -28.9255
18.8439 | -35.0197 | -28.9326
18.85 | -35.0202 | -28.9399
18.8561 | -35.0207 | -28.9467
18.8622 | -35.0211 | -28.9537
18.8683 | -35.0216 | -28.9605
18.8744 | -35.0222 | -28.9677
18.8805 | -35.0227 | -28.9751
18.8866 | -35.0232 | -28.9824
18.8927 | -35.0237 | -28.9894
18.8988 | -35.0242 | -28.9964
18.9049 | -35.0247 | -29.0033
18.911 | -35.0252 | -29.0105
18.9171 | -35.0258 | -29.0175
18.9232 | -35.0263 | -29.0241
18.9293 | -35.0267 | -29.0313
18.9354 | -35.0272 | -29.0386
18.9415 | -35.0277 | -29.0458
18.9476 | -35.0282 | -29.0526
18.9537 | -35.0286 | -29.0596
18.9598 | -35.0292 | -29.0661
18.9659 | -35.0297 | -29.073
18.972 | -35.0302 | -29.0799
18.9781 | -35.0306 | -29.0871
18.9843 | -35.0312 | -29.0943
18.9904 | -35.0316 | -29.1013
18.9965 | -35.0321 | -29.1082
19.0026 | -35.0326 | -29.1153
19.0087 | -35.0331 | -29.122
19.0148 | -35.0336 | -29.1286
19.0209 | -35.0341 | -29.1358
19.027 | -35.0346 | -29.1428
19.0331 | -35.035 | -29.1492
19.0392 | -35.0355 | -29.1567
19.0453 | -35.0361 | -29.1639
19.0514 | -35.0365 | -29.1707
19.0575 | -35.0371 | -29.178
19.0636 | -35.0374 | -29.1846
19.0697 | -35.038 | -29.1916
19.0758 | -35.0385 | -29.1985
19.0819 | -35.039 | -29.2056
19.088 | -35.0394 | -29.2125
19.0941 | -35.04 | -29.219
19.1002 | -35.0404 | -29.2256
19.1063 | -35.041 | -29.2327
19.1124 | -35.0415 | -29.2397
19.1185 | -35.042 | -29.2464
19.1246 | -35.0425 | -29.2535
19.1308 | -35.0431 | -29.2605
19.1369 | -35.0437 | -29.2671
19.143 | -35.0442 | -29.2741
19.1491 | -35.0447 | -29.2814
19.1552 | -35.0453 | -29.2889
19.1613 | -35.0457 | -29.2954
19.1674 | -35.0463 | -29.303
19.1735 | -35.0467 | -29.3098
19.1796 | -35.0471 | -29.3164
19.1857 | -35.0476 | -29.3232
19.1918 | -35.0481 | -29.3304
19.1979 | -35.0486 | -29.3376
19.204 | -35.0491 | -29.3445
19.2101 | -35.0497 | -29.3514
19.2162 | -35.0501 | -29.3583
19.2223 | -35.0506 | -29.3646
19.2284 | -35.0512 | -29.372
19.2345 | -35.0517 | -29.3793
19.2406 | -35.0522 | -29.3863
19.2467 | -35.0527 | -29.3928
19.2528 | -35.0532 | -29.3994
19.2589 | -35.0537 | -29.4058
19.265 | -35.0542 | -29.4129
19.2712 | -35.0547 | -29.4194
19.2773 | -35.0552 | -29.4266
19.2834 | -35.0557 | -29.433
19.2895 | -35.0562 | -29.4398
19.2956 | -35.0567 | -29.4472
19.3017 | -35.0572 | -29.4541
19.3078 | -35.0577 | -29.461
19.3139 | -35.0582 | -29.4679
19.32 | -35.0587 | -29.4747
19.3261 | -35.0592 | -29.4818
19.3322 | -35.0597 | -29.4885
19.3383 | -35.0602 | -29.4952
19.3444 | -35.0608 | -29.5021
19.3505 | -35.0612 | -29.5088
19.3566 | -35.0616 | -29.5154
19.3627 | -35.0621 | -29.5231
19.3688 | -35.0626 | -29.5298
19.3749 | -35.0631 | -29.5366
19.381 | -35.0635 | -29.5432
19.3871 | -35.064 | -29.5503
19.3932 | -35.0645 | -29.5572
19.3993 | -35.065 | -29.5638
19.4054 | -35.0655 | -29.5711
19.4115 | -35.066 | -29.578
19.4177 | -35.0665 | -29.5849
19.4238 | -35.067 | -29.5916
19.4299 | -35.0675 | -29.5985
19.436 | -35.068 | -29.6058
19.4421 | -35.0685 | -29.6121
19.4482 | -35.069 | -29.619
19.4543 | -35.0695 | -29.625
19.4604 | -35.07 | -29.632
19.4665 | -35.0705 | -29.6392
19.4726 | -35.071 | -29.6462
19.4787 | -35.0715 | -29.6533
19.4848 | -35.072 | -29.66
19.4909 | -35.0724 | -29.6668
19.497 | -35.0729 | -29.6734
19.5031 | -35.0734 | -29.6801
19.5092 | -35.0738 | -29.6868
19.5153 | -35.0744 | -29.6939
19.5214 | -35.0749 | -29.7007
19.5275 | -35.0753 | -29.7075
19.5336 | -35.0758 | -29.7142
19.5397 | -35.0763 | -29.721
19.5458 | -35.0768 | -29.7279
19.5519 | -35.0773 | -29.7349
19.5581 | -35.0778 | -29.7417
19.5642 | -35.0783 | -29.7487
19.5703 | -35.0788 | -29.7552
19.5764 | -35.0793 | -29.7623
19.5825 | -35.0798 | -29.7689
19.5886 | -35.0803 | -29.7759
19.5947 | -35.0808 | -29.7826
19.6008 | -35.0813 | -29.7894
19.6069 | -35.0818 | -29.7963
19.613 | -35.0823 | -29.8032
19.6191 | -35.0828 | -29.8098
19.6252 | -35.0833 | -29.8165
19.6313 | -35.0837 | -29.8233
19.6374 | -35.0842 | -29.8305
19.6435 | -35.0847 | -29.8373
19.6496 | -35.0851 | -29.844
19.6557 | -35.0856 | -29.8508
19.6618 | -35.0861 | -29.8574
19.6679 | -35.0866 | -29.8643
19.674 | -35.0871 | -29.8708
19.6801 | -35.0877 | -29.8776
19.6862 | -35.0882 | -29.8846
19.6923 | -35.0888 | -29.8909
19.6984 | -35.0893 | -29.897
19.7046 | -35.0898 | -29.9037
19.7107 | -35.0903 | -29.9104
19.7168 | -35.0908 | -29.9174
19.7229 | -35.0912 | -29.9243
19.729 | -35.0917 | -29.9314
19.7351 | -35.0922 | -29.9386
19.7412 | -35.0926 | -29.9454
19.7473 | -35.0931 | -29.9528
19.7534 | -35.0936 | -29.9599
19.7595 | -35.094 | -29.9671
19.7656 | -35.0945 | -29.9743
19.7717 | -35.0949 | -29.98
19.7778 | -35.0953 | -29.9873
19.7839 | -35.0958 | -29.9936
19.79 | -35.0962 | -30.0009
19.7961 | -35.0967 | -30.0079
19.8022 | -35.0973 | -30.0146
19.8083 | -35.0978 | -30.0215
19.8144 | -35.0983 | -30.0283
19.8205 | -35.0988 | -30.035
19.8266 | -35.0993 | -30.042
19.8327 | -35.0998 | -30.0486
19.8388 | -35.1003 | -30.0557
19.845 | -35.1008 | -30.0627
19.8511 | -35.1013 | -30.0702
19.8572 | -35.1017 | -30.0769
19.8633 | -35.1021 | -30.0837
19.8694 | -35.1028 | -30.0909
19.8755 | -35.1032 | -30.098
19.8816 | -35.1038 | -30.1045
19.8877 | -35.1043 | -30.1108
19.8938 | -35.1048 | -30.1172
19.8999 | -35.1053 | -30.1235
19.906 | -35.1059 | -30.1304
19.9121 | -35.1065 | -30.1373
19.9182 | -35.1069 | -30.1438
19.9243 | -35.1074 | -30.1507
19.9304 | -35.1079 | -30.1578
19.9365 | -35.1084 | -30.165
19.9426 | -35.1088 | -30.1714
19.9487 | -35.1094 | -30.1778
19.9548 | -35.1099 | -30.1851
19.9609 | -35.1103 | -30.1919
19.967 | -35.1108 | -30.199
19.9731 | -35.1114 | -30.2058
19.9792 | -35.1119 | -30.2122
19.9853 | -35.1125 | -30.2193
19.9915 | -35.113 | -30.2265
19.9976 | -35.1135 | -30.233
20.0037 | -35.1139 | -30.2399
20.0098 | -35.1144 | -30.2469
20.0159 | -35.115 | -30.2534
20.022 | -35.1155 | -30.2604
20.0281 | -35.116 | -30.2674
20.0342 | -35.1165 | -30.2739
20.0403 | -35.117 | -30.2801
20.0464 | -35.1175 | -30.2872
20.0525 | -35.118 | -30.2939
20.0586 | -35.1184 | -30.3001
20.0647 | -35.1189 | -30.3078
20.0708 | -35.1194 | -30.315
20.0769 | -35.1198 | -30.3223
20.083 | -35.1203 | -30.3276
20.0891 | -35.1208 | -30.3348
20.0952 | -35.1213 | -30.3417
20.1013 | -35.1218 | -30.3479
20.1074 | -35.1222 | -30.3549
20.1135 | -35.1227 | -30.3613
20.1196 | -35.1232 | -30.3683
20.1257 | -35.1237 | -30.3751
20.1319 | -35.1242 | -30.3817
20.138 | -35.1247 | -30.3882
20.1441 | -35.1251 | -30.3948
20.1502 | -35.1256 | -30.4017
20.1563 | -35.1262 | -30.4083
20.1624 | -35.1266 | -30.4145
20.1685 | -35.1271 | -30.4216
20.1746 | -35.1277 | -30.4283
20.1807 | -35.1282 | -30.435
20.1868 | -35.1287 | -30.4412
20.1929 | -35.1292 | -30.4479
20.199 | -35.1296 | -30.455
20.2051 | -35.1301 | -30.4615
20.2112 | -35.1306 | -30.4683
20.2173 | -35.1311 | -30.4749
20.2234 | -35.1316 | -30.4818
20.2295 | -35.1321 | -30.4883
20.2356 | -35.1326 | -30.4949
20.2417 | -35.1331 | -30.5014
20.2478 | -35.1335 | -30.5079
20.2539 | -35.1339 | -30.5146
20.26 | -35.1344 | -30.521
20.2661 | -35.135 | -30.5279
20.2723 | -35.1354 | -30.5344
20.2784 | -35.1358 | -30.5411
20.2845 | -35.1362 | -30.5478
20.2906 | -35.1368 | -30.5542
20.2967 | -35.1373 | -30.5598
20.3028 | -35.1378 | -30.5667
20.3089 | -35.1382 | -30.5733
20.315 | -35.1388 | -30.5798
20.3211 | -35.1392 | -30.5865
20.3272 | -35.1397 | -30.5932
20.3333 | -35.1402 | -30.5994
20.3394 | -35.1407 | -30.6057
20.3455 | -35.1412 | -30.6124
20.3516 | -35.1417 | -30.6188
20.3577 | -35.1421 | -30.6252
20.3638 | -35.1426 | -30.6315
20.3699 | -35.143 | -30.6378
20.376 | -35.1436 | -30.6446
20.3821 | -35.144 | -30.6508
20.3882 | -35.1444 | -30.6568
20.3943 | -35.1448 | -30.6639
20.4004 | -35.1453 | -30.6692
20.4065 | -35.1458 | -30.6749
20.4126 | -35.1463 | -30.681
20.4188 | -35.1469 | -30.6881
20.4249 | -35.1473 | -30.6948
20.431 | -35.1477 | -30.7014
20.4371 | -35.1482 | -30.7074
20.4432 | -35.1487 | -30.7142
20.4493 | -35.1492 | -30.7209
20.4554 | -35.1496 | -30.7278
20.4615 | -35.1501 | -30.7341
20.4676 | -35.1506 | -30.7401
20.4737 | -35.1511 | -30.7464
20.4798 | -35.1517 | -30.7528
20.4859 | -35.1522 | -30.7598
20.492 | -35.1528 | -30.7658
20.4981 | -35.1533 | -30.7715
20.5042 | -35.1536 | -30.7776
20.5103 | -35.1541 | -30.7842
20.5164 | -35.1545 | -30.7912
20.5225 | -35.155 | -30.798
20.5286 | -35.1555 | -30.8046
20.5347 | -35.156 | -30.8112
20.5408 | -35.1564 | -30.8178
20.5469 | -35.1569 | -30.824
20.553 | -35.1574 | -30.8311
20.5592 | -35.1578 | -30.8379
20.5653 | -35.1584 | -30.8451
20.5714 | -35.1588 | -30.8508
20.5775 | -35.1595 | -30.8571
20.5836 | -35.16 | -30.8635
20.5897 | -35.1607 | -30.8701
20.5958 | -35.1611 | -30.8761
20.6019 | -35.1615 | -30.8822
20.608 | -35.162 | -30.8885
20.6141 | -35.1625 | -30.8943
20.6202 | -35.163 | -30.9011
20.6263 | -35.1635 | -30.9076
20.6324 | -35.1641 | -30.9137
20.6385 | -35.1646 | -30.9198
20.6446 | -35.165 | -30.9273
20.6507 | -35.1654 | -30.934
20.6568 | -35.166 | -30.9403
20.6629 | -35.1664 | -30.9472
20.669 | -35.1669 | -30.9537
20.6751 | -35.1673 | -30.9603
20.6812 | -35.1678 | -30.9671
20.6873 | -35.1684 | -30.9736
20.6934 | -35.1689 | -30.9806
20.6995 | -35.1693 | -30.9876
20.7057 | -35.1697 | -30.9932
20.7118 | -35.1703 | -30.9999
20.7179 | -35.1709 | -31.005
20.724 | -35.1714 | -31.0115
20.7301 | -35.1718 | -31.0178
20.7362 | -35.1722 | -31.0242
20.7423 | -35.1727 | -31.0305
20.7484 | -35.1731 | -31.0369
20.7545 | -35.1736 | -31.0433
20.7606 | -35.1742 | -31.0496
20.7667 | -35.1746 | -31.0561
20.7728 | -35.1752 | -31.0625
20.7789 | -35.1756 | -31.0686
20.785 | -35.1761 | -31.0749
20.7911 | -35.1765 | -31.0815
20.7972 | -35.1769 | -31.0876
20.8033 | -35.1773 | -31.0942
20.8094 | -35.1779 | -31.1009
20.8155 | -35.1784 | -31.1073
20.8216 | -35.1789 | -31.114
20.8277 | -35.1793 | -31.1203
20.8338 | -35.1798 | -31.1268
20.8399 | -35.1803 | -31.1332
20.8461 | -35.1807 | -31.1395
20.8522 | -35.1812 | -31.1457
20.8583 | -35.1818 | -31.1515
20.8644 | -35.1823 | -31.1582
20.8705 | -35.1827 | -31.1648
20.8766 | -35.1832 | -31.171
20.8827 | -35.1838 | -31.1777
20.8888 | -35.1843 | -31.1835
20.8949 | -35.1847 | -31.1894
20.901 | -35.1852 | -31.1956
20.9071 | -35.1857 | -31.2014
20.9132 | -35.1861 | -31.2075
20.9193 | -35.1866 | -31.2138
20.9254 | -35.1871 | -31.2197
20.9315 | -35.1876 | -31.2256
20.9376 | -35.188 | -31.2321
20.9437 | -35.1885 | -31.2387
20.9498 | -35.189 | -31.2449
20.9559 | -35.1895 | -31.2512
20.962 | -35.19 | -31.2578
20.9681 | -35.1904 | -31.2647
20.9742 | -35.191 | -31.2707
20.9803 | -35.1915 | -31.2763
20.9864 | -35.1919 | -31.283
20.9926 | -35.1924 | -31.2894
20.9987 | -35.193 | -31.295
21.0048 | -35.1934 | -31.3009
21.0109 | -35.1938 | -31.3071
21.017 | -35.1943 | -31.3137
21.0231 | -35.1948 | -31.3202
21.0292 | -35.1953 | -31.3269
21.0353 | -35.1958 | -31.3329
21.0414 | -35.1964 | -31.3389
21.0475 | -35.1968 | -31.3456
21.0536 | -35.1973 | -31.3519
21.0597 | -35.1979 | -31.3578
21.0658 | -35.1983 | -31.364
21.0719 | -35.1989 | -31.3708
21.078 | -35.1994 | -31.3775
21.0841 | -35.1998 | -31.3836
21.0902 | -35.2003 | -31.3897
21.0963 | -35.201 | -31.3959
21.1024 | -35.2015 | -31.4021
21.1085 | -35.2018 | -31.4085
21.1146 | -35.2023 | -31.4139
21.1207 | -35.2029 | -31.4205
21.1268 | -35.2033 | -31.4274
21.133 | -35.2039 | -31.434
21.1391 | -35.2044 | -31.44
21.1452 | -35.2049 | -31.4466
21.1513 | -35.2053 | -31.453
21.1574 | -35.2059 | -31.4593
21.1635 | -35.2064 | -31.4656
21.1696 | -35.2068 | -31.4723
21.1757 | -35.2074 | -31.4782
21.1818 | -35.2078 | -31.4847
21.1879 | -35.2082 | -31.4907
21.194 | -35.2089 | -31.4969
21.2001 | -35.2095 | -31.5032
21.2062 | -35.2099 | -31.5092
21.2123 | -35.2103 | -31.5144
21.2184 | -35.2107 | -31.5198
21.2245 | -35.2111 | -31.5265
21.2306 | -35.2116 | -31.5325
21.2367 | -35.2121 | -31.5385
21.2428 | -35.2126 | -31.5449
21.2489 | -35.213 | -31.5504
21.255 | -35.2135 | -31.5563
21.2611 | -35.214 | -31.5621
21.2672 | -35.2145 | -31.5679
21.2733 | -35.2148 | -31.5745
21.2795 | -35.2154 | -31.5794
21.2856 | -35.2159 | -31.5855
21.2917 | -35.2164 | -31.5914
21.2978 | -35.2169 | -31.598
21.3039 | -35.2173 | -31.6037
21.31 | -35.2178 | -31.6097
21.3161 | -35.2184 | -31.6154
21.3222 | -35.2189 | -31.6212
21.3283 | -35.2194 | -31.6278
21.3344 | -35.2199 | -31.634
21.3405 | -35.2203 | -31.6406
21.3466 | -35.2209 | -31.6471
21.3527 | -35.2214 | -31.6526
21.3588 | -35.222 | -31.6584
21.3649 | -35.2225 | -31.6643
21.371 | -35.223 | -31.6702
21.3771 | -35.2235 | -31.6763
21.3832 | -35.224 | -31.6822
21.3893 | -35.2245 | -31.6885
21.3954 | -35.225 | -31.6947
21.4015 | -35.2256 | -31.7008
21.4076 | -35.226 | -31.7066
21.4137 | -35.2265 | -31.713
21.4199 | -35.227 | -31.7189
21.426 | -35.2275 | -31.7237
21.4321 | -35.2279 | -31.7298
21.4382 | -35.2284 | -31.7359
21.4443 | -35.2289 | -31.7416
21.4504 | -35.2293 | -31.7479
21.4565 | -35.2297 | -31.7538
21.4626 | -35.2302 | -31.7599
21.4687 | -35.2307 | -31.766
21.4748 | -35.2311 | -31.7725
21.4809 | -35.2316 | -31.7785
21.487 | -35.2321 | -31.7845
21.4931 | -35.2327 | -31.7905
21.4992 | -35.2332 | -31.7965
21.5053 | -35.2336 | -31.8028
21.5114 | -35.2341 | -31.8091
21.5175 | -35.2345 | -31.8153
21.5236 | -35.2349 | -31.8214
21.5297 | -35.2354 | -31.8273
21.5358 | -35.2359 | -31.8335
21.5419 | -35.2364 | -31.8404
21.548 | -35.2369 | -31.8462
21.5541 | -35.2374 | -31.8516
21.5602 | -35.2379 | -31.8571
21.5664 | -35.2384 | -31.8624
21.5725 | -35.2389 | -31.8682
21.5786 | -35.2393 | -31.8749
21.5847 | -35.2398 | -31.8809
21.5908 | -35.2403 | -31.887
21.5969 | -35.2408 | -31.8925
21.603 | -35.2412 | -31.8984
21.6091 | -35.2417 | -31.9046
21.6152 | -35.2422 | -31.9111
21.6213 | -35.2426 | -31.9178
21.6274 | -35.2431 | -31.9232
21.6335 | -35.2436 | -31.9292
21.6396 | -35.2442 | -31.9349
21.6457 | -35.2446 | -31.941
21.6518 | -35.2451 | -31.947
21.6579 | -35.2456 | -31.9527
21.664 | -35.2461 | -31.9593
21.6701 | -35.2466 | -31.9646
21.6762 | -35.2471 | -31.9708
21.6823 | -35.2475 | -31.9771
21.6884 | -35.248 | -31.9836
21.6945 | -35.2484 | -31.9891
21.7006 | -35.2489 | -31.9957
21.7068 | -35.2494 | -32.0018
21.7129 | -35.2499 | -32.0074
21.719 | -35.2504 | -32.013
21.7251 | -35.2509 | -32.0196
21.7312 | -35.2515 | -32.0253
21.7373 | -35.2519 | -32.0309
21.7434 | -35.2524 | -32.0369
21.7495 | -35.2528 | -32.0429
21.7556 | -35.2533 | -32.0489
21.7617 | -35.2538 | -32.0554
21.7678 | -35.2544 | -32.0617
21.7739 | -35.2549 | -32.0674
21.78 | -35.2554 | -32.0727
21.7861 | -35.256 | -32.0783
21.7922 | -35.2565 | -32.0842
21.7983 | -35.2569 | -32.09
21.8044 | -35.2574 | -32.0961
21.8105 | -35.2579 | -32.1021
21.8166 | -35.2584 | -32.1085
21.8227 | -35.2588 | -32.1142
21.8288 | -35.2594 | -32.1194
21.8349 | -35.2598 | -32.125
21.841 | -35.2604 | -32.1305
21.8471 | -35.2608 | -32.1361
21.8533 | -35.2613 | -32.1419
21.8594 | -35.2618 | -32.1483
21.8655 | -35.2622 | -32.154
21.8716 | -35.2627 | -32.1603
21.8777 | -35.2633 | -32.1655
21.8838 | -35.2637 | -32.1714
21.8899 | -35.2641 | -32.1774
21.896 | -35.2647 | -32.1835
21.9021 | -35.2651 | -32.1894
21.9082 | -35.2655 | -32.1951
21.9143 | -35.266 | -32.2008
21.9204 | -35.2666 | -32.2068
21.9265 | -35.2671 | -32.2125
21.9326 | -35.2675 | -32.2184
21.9387 | -35.2681 | -32.2248
21.9448 | -35.2686 | -32.2313
21.9509 | -35.2692 | -32.2365
21.957 | -35.2697 | -32.2426
21.9631 | -35.2702 | -32.2487
21.9692 | -35.2707 | -32.2546
21.9753 | -35.2711 | -32.2604
21.9814 | -35.2716 | -32.2662
21.9875 | -35.2721 | -32.271
21.9937 | -35.2725 | -32.2773
21.9998 | -35.2731 | -32.2837
22.0059 | -35.2735 | -32.2897
22.012 | -35.2741 | -32.2956
22.0181 | -35.2747 | -32.3014
22.0242 | -35.2752 | -32.3073
22.0303 | -35.2757 | -32.3137
22.0364 | -35.2762 | -32.3194
22.0425 | -35.2766 | -32.325
22.0486 | -35.277 | -32.331
22.0547 | -35.2775 | -32.3364
22.0608 | -35.2779 | -32.3422
22.0669 | -35.2784 | -32.348
22.073 | -35.2789 | -32.3536
22.0791 | -35.2794 | -32.3595
22.0852 | -35.2798 | -32.3656
22.0913 | -35.2804 | -32.3722
22.0974 | -35.2808 | -32.378
22.1035 | -35.2812 | -32.3845
22.1096 | -35.2817 | -32.391
22.1157 | -35.2823 | -32.3964
22.1218 | -35.2828 | -32.4028
22.1279 | -35.2832 | -32.4081
22.134 | -35.2837 | -32.4143
22.1402 | -35.2842 | -32.4202
22.1463 | -35.2846 | -32.4257
22.1524 | -35.2851 | -32.4316
22.1585 | -35.2856 | -32.438
22.1646 | -35.2861 | -32.443
22.1707 | -35.2866 | -32.4494
22.1768 | -35.2872 | -32.4553
22.1829 | -35.2878 | -32.4611
22.189 | -35.2883 | -32.467
22.1951 | -35.2889 | -32.4724
22.2012 | -35.2893 | -32.4782
22.2073 | -35.2898 | -32.4843
22.2134 | -35.2903 | -32.4902
22.2195 | -35.2909 | -32.4962
22.2256 | -35.2912 | -32.5021
22.2317 | -35.2917 | -32.5078
22.2378 | -35.2921 | -32.5132
22.2439 | -35.2927 | -32.5196
22.25 | -35.2931 | -32.5249
22.2561 | -35.2936 | -32.5305
22.2622 | -35.2942 | -32.5368
22.2683 | -35.2946 | -32.5422
22.2744 | -35.2951 | -32.5484
22.2806 | -35.2955 | -32.5545
22.2867 | -35.296 | -32.5597
22.2928 | -35.2965 | -32.5658
22.2989 | -35.297 | -32.5717
22.305 | -35.2976 | -32.5777
22.3111 | -35.298 | -32.5839
22.3172 | -35.2985 | -32.5892
22.3233 | -35.299 | -32.5954
22.3294 | -35.2994 | -32.601
22.3355 | -35.2998 | -32.6073
22.3416 | -35.3004 | -32.6132
22.3477 | -35.3009 | -32.6192
22.3538 | -35.3013 | -32.6253
22.3599 | -35.3018 | -32.6312
22.366 | -35.3022 | -32.6369
22.3721 | -35.3027 | -32.6423
22.3782 | -35.3032 | -32.6478
22.3843 | -35.3038 | -32.6534
22.3904 | -35.3043 | -32.6592
22.3965 | -35.3048 | -32.665
22.4026 | -35.3052 | -32.6704
22.4087 | -35.3057 | -32.6766
22.4148 | -35.3061 | -32.6821
22.4209 | -35.3066 | -32.688
22.4271 | -35.3071 | -32.6932
22.4332 | -35.3075 | -32.6986
22.4393 | -35.308 | -32.7041
22.4454 | -35.3084 | -32.7099
22.4515 | -35.3088 | -32.7158
22.4576 | -35.3093 | -32.7213
22.4637 | -35.3097 | -32.7274
22.4698 | -35.3102 | -32.7334
22.4759 | -35.3107 | -32.7394
22.482 | -35.3111 | -32.7448
22.4881 | -35.3115 | -32.7502
22.4942 | -35.3119 | -32.7564
22.5003 | -35.3124 | -32.7626
22.5064 | -35.3129 | -32.7677
22.5125 | -35.3132 | -32.774
22.5186 | -35.3136 | -32.7791
22.5247 | -35.3141 | -32.7855
22.5308 | -35.3146 | -32.7917
22.5369 | -35.3151 | -32.7981
22.543 | -35.3156 | -32.8042
22.5491 | -35.3159 | -32.8101
22.5552 | -35.3164 | -32.8157
22.5613 | -35.3169 | -32.8215
22.5675 | -35.3174 | -32.8276
22.5736 | -35.3179 | -32.8343
22.5797 | -35.3183 | -32.8404
22.5858 | -35.3186 | -32.8454
22.5919 | -35.3191 | -32.8506
22.598 | -35.3196 | -32.8558
22.6041 | -35.3201 | -32.8618
22.6102 | -35.3206 | -32.8672
22.6163 | -35.321 | -32.8732
22.6224 | -35.3215 | -32.8793
22.6285 | -35.3219 | -32.8858
22.6346 | -35.3223 | -32.8916
22.6407 | -35.3227 | -32.8976
22.6468 | -35.3232 | -32.9034
22.6529 | -35.3235 | -32.9088
22.659 | -35.324 | -32.9146
22.6651 | -35.3245 | -32.92
22.6712 | -35.3249 | -32.926
22.6773 | -35.3255 | -32.9316
22.6834 | -35.3259 | -32.9372
22.6895 | -35.3263 | -32.9436
22.6956 | -35.3269 | -32.9492
22.7017 | -35.3273 | -32.9555
22.7079 | -35.3278 | -32.9619
22.714 | -35.3281 | -32.9675
22.7201 | -35.3286 | -32.9733
22.7262 | -35.329 | -32.9788
22.7323 | -35.3295 | -32.9843
22.7384 | -35.3299 | -32.9903
22.7445 | -35.3303 | -32.9965
22.7506 | -35.3308 | -33.0025
22.7567 | -35.3312 | -33.0086
22.7628 | -35.3316 | -33.0142
22.7689 | -35.3321 | -33.0197
22.775 | -35.3326 | -33.0255
22.7811 | -35.3329 | -33.0311
22.7872 | -35.3334 | -33.0369
22.7933 | -35.3338 | -33.0426
22.7994 | -35.3343 | -33.0483
22.8055 | -35.3348 | -33.054
22.8116 | -35.3353 | -33.06
22.8177 | -35.3357 | -33.0659
22.8238 | -35.3362 | -33.072
22.8299 | -35.3366 | -33.0775
22.836 | -35.3372 | -33.0839
22.8421 | -35.3377 | -33.0896
22.8482 | -35.3382 | -33.0957
22.8544 | -35.3387 | -33.1018
22.8605 | -35.3392 | -33.1075
22.8666 | -35.3396 | -33.1137
22.8727 | -35.3401 | -33.1195
22.8788 | -35.3405 | -33.1254
22.8849 | -35.341 | -33.1309
22.891 | -35.3415 | -33.1368
22.8971 | -35.342 | -33.1425
22.9032 | -35.3424 | -33.1481
22.9093 | -35.3429 | -33.1541
22.9154 | -35.3435 | -33.1598
22.9215 | -35.344 | -33.1653
22.9276 | -35.3444 | -33.1711
22.9337 | -35.3449 | -33.1763
22.9398 | -35.3454 | -33.1824
22.9459 | -35.3459 | -33.1883
22.952 | -35.3464 | -33.1937
22.9581 | -35.3468 | -33.1996
22.9642 | -35.3473 | -33.2054
22.9703 | -35.3477 | -33.2112
22.9764 | -35.3483 | -33.2168
22.9825 | -35.3487 | -33.2224
22.9886 | -35.3491 | -33.2281
22.9948 | -35.3496 | -33.2339
23.0009 | -35.35 | -33.2398
23.007 | -35.3504 | -33.2455
23.0131 | -35.3509 | -33.2518
23.0192 | -35.3514 | -33.2573
23.0253 | -35.3518 | -33.2628
23.0314 | -35.3523 | -33.2683
23.0375 | -35.3529 | -33.2739
23.0436 | -35.3533 | -33.28
23.0497 | -35.3537 | -33.2856
23.0558 | -35.3542 | -33.2912
23.0619 | -35.3546 | -33.2975
23.068 | -35.3551 | -33.3032
23.0741 | -35.3555 | -33.3088
23.0802 | -35.3561 | -33.3147
23.0863 | -35.3564 | -33.3204
23.0924 | -35.3569 | -33.3266
23.0985 | -35.3574 | -33.3321
23.1046 | -35.3579 | -33.3375
23.1107 | -35.3582 | -33.3433
23.1168 | -35.3586 | -33.3491
23.1229 | -35.3591 | -33.3553
23.129 | -35.3595 | -33.3606
23.1351 | -35.36 | -33.3659
23.1413 | -35.3604 | -33.3717
23.1474 | -35.3607 | -33.3773
23.1535 | -35.3611 | -33.3828
23.1596 | -35.3615 | -33.3886
23.1657 | -35.362 | -33.3942
23.1718 | -35.3624 | -33.4001
23.1779 | -35.3629 | -33.4062
23.184 | -35.3632 | -33.4121
23.1901 | -35.3637 | -33.4178
23.1962 | -35.3641 | -33.4231
23.2023 | -35.3645 | -33.4291
23.2084 | -35.365 | -33.4351
23.2145 | -35.3654 | -33.4399
23.2206 | -35.3658 | -33.4454
23.2267 | -35.3663 | -33.4516
23.2328 | -35.3666 | -33.4566
23.2389 | -35.3669 | -33.4623
23.245 | -35.3673 | -33.4681
23.2511 | -35.3677 | -33.4741
23.2572 | -35.368 | -33.4791
23.2633 | -35.3685 | -33.4849
23.2694 | -35.3689 | -33.4903
23.2755 | -35.3694 | -33.4963
23.2817 | -35.3698 | -33.502
23.2878 | -35.3702 | -33.508
23.2939 | -35.3707 | -33.5141
23.3 | -35.3711 | -33.52
23.3061 | -35.3715 | -33.5256
23.3122 | -35.372 | -33.5314
23.3183 | -35.3724 | -33.5371
23.3244 | -35.3729 | -33.5429
23.3305 | -35.3734 | -33.5495
23.3366 | -35.3738 | -33.5555
23.3427 | -35.3742 | -33.5613
23.3488 | -35.3747 | -33.5669
23.3549 | -35.3752 | -33.5725
23.361 | -35.3756 | -33.5784
23.3671 | -35.3761 | -33.584
23.3732 | -35.3766 | -33.5903
23.3793 | -35.3771 | -33.5958
23.3854 | -35.3775 | -33.6023
23.3915 | -35.3779 | -33.6074
23.3976 | -35.3783 | -33.6129
23.4037 | -35.3786 | -33.6189
23.4098 | -35.3791 | -33.6256
23.4159 | -35.3796 | -33.6312
23.422 | -35.3801 | -33.6369
23.4282 | -35.3805 | -33.643
23.4343 | -35.381 | -33.6493
23.4404 | -35.3814 | -33.6553
23.4465 | -35.3818 | -33.6606
23.4526 | -35.3823 | -33.6668
23.4587 | -35.3827 | -33.6726
23.4648 | -35.3831 | -33.6786
23.4709 | -35.3835 | -33.684
23.477 | -35.3839 | -33.6901
23.4831 | -35.3844 | -33.6959
23.4892 | -35.3849 | -33.7016
23.4953 | -35.3854 | -33.707
23.5014 | -35.3858 | -33.7118
23.5075 | -35.3863 | -33.717
23.5136 | -35.3867 | -33.7227
23.5197 | -35.3871 | -33.7291
23.5258 | -35.3875 | -33.7348
23.5319 | -35.3879 | -33.741
23.538 | -35.3883 | -33.7472
23.5441 | -35.3887 | -33.7531
23.5502 | -35.3891 | -33.7592
23.5563 | -35.3895 | -33.7651
23.5624 | -35.3899 | -33.7712
23.5686 | -35.3903 | -33.7774
23.5747 | -35.3907 | -33.7828
23.5808 | -35.3913 | -33.7888
23.5869 | -35.3917 | -33.7951
23.593 | -35.3921 | -33.8006
23.5991 | -35.3926 | -33.8062
23.6052 | -35.393 | -33.8121
23.6113 | -35.3934 | -33.8176
23.6174 | -35.3939 | -33.8235
23.6235 | -35.3943 | -33.829
23.6296 | -35.3947 | -33.8351
23.6357 | -35.3951 | -33.8408
23.6418 | -35.3954 | -33.8461
23.6479 | -35.3959 | -33.8526
23.654 | -35.3963 | -33.8586
23.6601 | -35.3967 | -33.8646
23.6662 | -35.3971 | -33.8706
23.6723 | -35.3975 | -33.8762
23.6784 | -35.3979 | -33.8821
23.6845 | -35.3983 | -33.8876
23.6906 | -35.3987 | -33.8937
23.6967 | -35.3991 | -33.8993
23.7028 | -35.3995 | -33.9051
23.7089 | -35.3998 | -33.9113
23.7151 | -35.4002 | -33.9171
23.7212 | -35.4006 | -33.9227
23.7273 | -35.401 | -33.9287
23.7334 | -35.4014 | -33.9348
23.7395 | -35.4018 | -33.9403
23.7456 | -35.4023 | -33.946
23.7517 | -35.4026 | -33.9522
23.7578 | -35.403 | -33.958
23.7639 | -35.4035 | -33.9637
23.77 | -35.4039 | -33.9698
23.7761 | -35.4043 | -33.976
23.7822 | -35.4047 | -33.9819
23.7883 | -35.4051 | -33.9876
23.7944 | -35.4057 | -33.9933
23.8005 | -35.4061 | -33.9993
23.8066 | -35.4065 | -34.005
23.8127 | -35.4069 | -34.0097
23.8188 | -35.4073 | -34.0153
23.8249 | -35.4076 | -34.0213
23.831 | -35.408 | -34.027
23.8371 | -35.4085 | -34.0323
23.8432 | -35.4089 | -34.0379
23.8493 | -35.4092 | -34.0436
23.8555 | -35.4097 | -34.0492
23.8616 | -35.4102 | -34.0551
23.8677 | -35.4105 | -34.0607
23.8738 | -35.4109 | -34.0657
23.8799 | -35.4112 | -34.0713
23.886 | -35.4117 | -34.0775
23.8921 | -35.4121 | -34.0831
23.8982 | -35.4125 | -34.0895
23.9043 | -35.4128 | -34.0955
23.9104 | -35.4131 | -34.1008
23.9165 | -35.4135 | -34.1059
23.9226 | -35.4139 | -34.1111
23.9287 | -35.4143 | -34.1177
23.9348 | -35.4146 | -34.123
23.9409 | -35.415 | -34.1292
23.947 | -35.4154 | -34.1347
23.9531 | -35.4158 | -34.1395
23.9592 | -35.4161 | -34.1455
23.9653 | -35.4166 | -34.1512
23.9714 | -35.417 | -34.157
23.9775 | -35.4174 | -34.1629
23.9836 | -35.4178 | -34.1696
23.9897 | -35.4181 | -34.176
23.9958 | -35.4185 | -34.1822
24.002 | -35.4188 | -34.1875
24.0081 | -35.4192 | -34.1929
24.0142 | -35.4195 | -34.1981
24.0203 | -35.4198 | -34.204
24.0264 | -35.4202 | -34.2104
24.0325 | -35.4206 | -34.2161
24.0386 | -35.4211 | -34.2218
24.0447 | -35.4214 | -34.2279
24.0508 | -35.4217 | -34.2341
24.0569 | -35.4222 | -34.2397
24.063 | -35.4225 | -34.2453
24.0691 | -35.423 | -34.2509
24.0752 | -35.4233 | -34.2566
24.0813 | -35.4237 | -34.2624
24.0874 | -35.4242 | -34.2679
24.0935 | -35.4245 | -34.2735
24.0996 | -35.4246 | -34.2791
24.1057 | -35.4249 | -34.285
24.1118 | -35.4253 | -34.2913
24.1179 | -35.4257 | -34.2973
24.124 | -35.4261 | -34.3033
24.1301 | -35.4265 | -34.3091
24.1362 | -35.4269 | -34.315
24.1424 | -35.4273 | -34.3202
24.1485 | -35.4276 | -34.3257
24.1546 | -35.4279 | -34.3318
24.1607 | -35.4283 | -34.337
24.1668 | -35.4287 | -34.3432
24.1729 | -35.429 | -34.3488
24.179 | -35.4294 | -34.3547
24.1851 | -35.4297 | -34.3609
24.1912 | -35.4301 | -34.3671
24.1973 | -35.4304 | -34.3727
24.2034 | -35.4308 | -34.3777
24.2095 | -35.4312 | -34.3836
24.2156 | -35.4316 | -34.3892
24.2217 | -35.432 | -34.3953
24.2278 | -35.4323 | -34.4005
24.2339 | -35.4327 | -34.4066
24.24 | -35.4329 | -34.4136
24.2461 | -35.4333 | -34.4195
24.2522 | -35.4335 | -34.425
24.2583 | -35.4339 | -34.4314
24.2644 | -35.4342 | -34.4373
24.2705 | -35.4345 | -34.4427
24.2766 | -35.4349 | -34.4483
24.2827 | -35.4353 | -34.4538
24.2889 | -35.4358 | -34.4596
24.295 | -35.4362 | -34.4655
24.3011 | -35.4366 | -34.4715
24.3072 | -35.437 | -34.4774
24.3133 | -35.4374 | -34.4832
24.3194 | -35.4377 | -34.4886
24.3255 | -35.4381 | -34.4944
24.3316 | -35.4384 | -34.5002
24.3377 | -35.4387 | -34.5059
24.3438 | -35.439 | -34.5105
24.3499 | -35.4394 | -34.5168
24.356 | -35.4397 | -34.5224
24.3621 | -35.4401 | -34.5283
24.3682 | -35.4404 | -34.5339
24.3743 | -35.4408 | -34.5402
24.3804 | -35.4412 | -34.5461
24.3865 | -35.4415 | -34.5524
24.3926 | -35.4418 | -34.5583
24.3987 | -35.4421 | -34.5642
24.4048 | -35.4425 | -34.57
24.4109 | -35.4427 | -34.5765
24.417 | -35.443 | -34.5824
24.4231 | -35.4433 | -34.5885
24.4293 | -35.4437 | -34.5945
24.4354 | -35.444 | -34.6005
24.4415 | -35.4444 | -34.606
24.4476 | -35.4448 | -34.6118
24.4537 | -35.4452 | -34.6174
24.4598 | -35.4455 | -34.6236
24.4659 | -35.4459 | -34.6297
24.472 | -35.4462 | -34.6352
24.4781 | -35.4465 | -34.6412
24.4842 | -35.4469 | -34.6463
24.4903 | -35.4472 | -34.6519
24.4964 | -35.4476 | -34.6573
24.5025 | -35.448 | -34.6633
24.5086 | -35.4484 | -34.6686
24.5147 | -35.4488 | -34.6738
24.5208 | -35.449 | -34.6791
24.5269 | -35.4494 | -34.6845
24.533 | -35.4497 | -34.6899
24.5391 | -35.4502 | -34.6957
24.5452 | -35.4505 | -34.7018
24.5513 | -35.4509 | -34.7072
24.5574 | -35.4512 | -34.7133
24.5635 | -35.4516 | -34.7186
24.5696 | -35.452 | -34.7245
24.5758 | -35.4524 | -34.7301
24.5819 | -35.4527 | -34.7357
24.588 | -35.453 | -34.7419
24.5941 | -35.4534 | -34.7468
24.6002 | -35.4537 | -34.7532
24.6063 | -35.454 | -34.759
24.6124 | -35.4543 | -34.7649
24.6185 | -35.4547 | -34.7715
24.6246 | -35.455 | -34.7772
24.6307 | -35.4553 | -34.7831
24.6368 | -35.4556 | -34.7887
24.6429 | -35.456 | -34.7946
24.649 | -35.4564 | -34.8
24.6551 | -35.4567 | -34.8056
24.6612 | -35.457 | -34.8113
24.6673 | -35.4574 | -34.8169
24.6734 | -35.4578 | -34.8231
24.6795 | -35.4581 | -34.8287
24.6856 | -35.4585 | -34.8344
24.6917 | -35.4588 | -34.8393
24.6978 | -35.4592 | -34.8453
24.7039 | -35.4596 | -34.8511
24.71 | -35.4599 | -34.8571
24.7162 | -35.4602 | -34.8626
24.7223 | -35.4605 | -34.8681
24.7284 | -35.4608 | -34.8736
24.7345 | -35.4612 | -34.8793
24.7406 | -35.4615 | -34.8856
24.7467 | -35.4619 | -34.8923
24.7528 | -35.4622 | -34.8976
24.7589 | -35.4625 | -34.9032
24.765 | -35.4627 | -34.9083
24.7711 | -35.4631 | -34.9145
24.7772 | -35.4633 | -34.9206
24.7833 | -35.4637 | -34.926
24.7894 | -35.4641 | -34.9314
24.7955 | -35.4644 | -34.9375
24.8016 | -35.4648 | -34.9431
24.8077 | -35.4652 | -34.9495
24.8138 | -35.4655 | -34.9551
24.8199 | -35.4659 | -34.9605
24.826 | -35.4662 | -34.9664
24.8321 | -35.4665 | -34.9724
24.8382 | -35.4669 | -34.9782
24.8443 | -35.4672 | -34.984
24.8504 | -35.4676 | -34.9891
24.8565 | -35.4679 | -34.9946
24.8627 | -35.4682 | -35.0005
24.8688 | -35.4685 | -35.0059
24.8749 | -35.4689 | -35.0114
24.881 | -35.4692 | -35.0169
24.8871 | -35.4695 | -35.0228
24.8932 | -35.4699 | -35.0279
24.8993 | -35.4703 | -35.0331
24.9054 | -35.4706 | -35.0394
24.9115 | -35.4709 | -35.0448
24.9176 | -35.4713 | -35.0511
24.9237 | -35.4716 | -35.0572
24.9298 | -35.472 | -35.0635
24.9359 | -35.4723 | -35.0697
24.942 | -35.4725 | -35.0752
24.9481 | -35.4729 | -35.0809
24.9542 | -35.4732 | -35.0862
24.9603 | -35.4735 | -35.0919
24.9664 | -35.4739 | -35.0975
24.9725 | -35.4742 | -35.1029
24.9786 | -35.4745 | -35.1087
24.9847 | -35.4748 | -35.1142
24.9908 | -35.4751 | -35.1202
24.9969 | -35.4753 | -35.1259
25.0031 | -35.4756 | -35.1319
25.0092 | -35.4759 | -35.1374
25.0153 | -35.4763 | -35.1429
25.0214 | -35.4766 | -35.1485
25.0275 | -35.477 | -35.1541
25.0336 | -35.4772 | -35.1602
25.0397 | -35.4776 | -35.1654
25.0458 | -35.4779 | -35.1707
25.0519 | -35.4782 | -35.1769
25.058 | -35.4786 | -35.1826
25.0641 | -35.4788 | -35.1884
25.0702 | -35.4791 | -35.1937
25.0763 | -35.4794 | -35.1993
25.0824 | -35.4796 | -35.2049
25.0885 | -35.48 | -35.2101
25.0946 | -35.4803 | -35.2163
25.1007 | -35.4807 | -35.2218
25.1068 | -35.4811 | -35.2273
25.1129 | -35.4814 | -35.2331
25.119 | -35.4818 | -35.2389
25.1251 | -35.4821 | -35.2445
25.1312 | -35.4824 | -35.2504
25.1373 | -35.4828 | -35.2565
25.1435 | -35.4831 | -35.2624
25.1496 | -35.4834 | -35.2678
25.1557 | -35.4836 | -35.2736
25.1618 | -35.4839 | -35.2793
25.1679 | -35.4842 | -35.2848
25.174 | -35.4846 | -35.2905
25.1801 | -35.4849 | -35.2962
25.1862 | -35.4854 | -35.3021
25.1923 | -35.4857 | -35.3079
25.1984 | -35.4861 | -35.3135
25.2045 | -35.4865 | -35.3191
25.2106 | -35.4868 | -35.3253
25.2167 | -35.4871 | -35.3312
25.2228 | -35.4874 | -35.3369
25.2289 | -35.4877 | -35.3421
25.235 | -35.488 | -35.3478
25.2411 | -35.4883 | -35.3535
25.2472 | -35.4886 | -35.3594
25.2533 | -35.489 | -35.3651
25.2594 | -35.4892 | -35.3705
25.2655 | -35.4895 | -35.3765
25.2716 | -35.4899 | -35.3821
25.2777 | -35.4902 | -35.3883
25.2838 | -35.4906 | -35.3941
25.29 | -35.4909 | -35.3993
25.2961 | -35.4913 | -35.4049
25.3022 | -35.4916 | -35.4111
25.3083 | -35.4919 | -35.4168
25.3144 | -35.4922 | -35.4224
25.3205 | -35.4925 | -35.4276
25.3266 | -35.4928 | -35.4332
25.3327 | -35.4931 | -35.4387
25.3388 | -35.4934 | -35.4445
25.3449 | -35.4938 | -35.4505
25.351 | -35.4941 | -35.4567
25.3571 | -35.4944 | -35.4627
25.3632 | -35.4947 | -35.4681
25.3693 | -35.4949 | -35.4739
25.3754 | -35.4952 | -35.48
25.3815 | -35.4955 | -35.4856
25.3876 | -35.4959 | -35.4911
25.3937 | -35.4962 | -35.4969
25.3998 | -35.4965 | -35.5024
25.4059 | -35.4968 | -35.5082
25.412 | -35.4971 | -35.5136
25.4181 | -35.4974 | -35.5193
25.4242 | -35.4978 | -35.525
25.4304 | -35.4981 | -35.5307
25.4365 | -35.4984 | -35.5365
25.4426 | -35.4987 | -35.542
25.4487 | -35.499 | -35.5476
25.4548 | -35.4992 | -35.5533
25.4609 | -35.4996 | -35.5587
25.467 | -35.4999 | -35.5641
25.4731 | -35.5003 | -35.5698
25.4792 | -35.5006 | -35.5754
25.4853 | -35.5009 | -35.5806
25.4914 | -35.5012 | -35.5861
25.4975 | -35.5015 | -35.5917
25.5036 | -35.5017 | -35.5974
25.5097 | -35.5019 | -35.6028
25.5158 | -35.5022 | -35.6083
25.5219 | -35.5025 | -35.614
25.528 | -35.5028 | -35.6196
25.5341 | -35.5031 | -35.6251
25.5402 | -35.5034 | -35.631
25.5463 | -35.5037 | -35.6371
25.5524 | -35.504 | -35.643
25.5585 | -35.5043 | -35.6481
25.5646 | -35.5046 | -35.6536
25.5707 | -35.505 | -35.6594
25.5769 | -35.5053 | -35.6644
25.583 | -35.5056 | -35.6699
25.5891 | -35.506 | -35.6758
25.5952 | -35.5063 | -35.6812
25.6013 | -35.5066 | -35.687
25.6074 | -35.5069 | -35.6923
25.6135 | -35.5072 | -35.6975
25.6196 | -35.5075 | -35.7033
25.6257 | -35.5078 | -35.7088
25.6318 | -35.508 | -35.7141
25.6379 | -35.5084 | -35.72
25.644 | -35.5086 | -35.7252
25.6501 | -35.5089 | -35.7307
25.6562 | -35.5092 | -35.7364
25.6623 | -35.5095 | -35.7422
25.6684 | -35.5098 | -35.7477
25.6745 | -35.5101 | -35.753
25.6806 | -35.5104 | -35.7588
25.6867 | -35.5107 | -35.7641
25.6928 | -35.511 | -35.7696
25.6989 | -35.5113 | -35.7756
25.705 | -35.5115 | -35.7818
25.7111 | -35.5118 | -35.7871
25.7173 | -35.5121 | -35.7927
25.7234 | -35.5125 | -35.7984
25.7295 | -35.5128 | -35.804
25.7356 | -35.5131 | -35.8097
25.7417 | -35.5134 | -35.8152
25.7478 | -35.5138 | -35.8206
25.7539 | -35.5141 | -35.8259
25.76 | -35.5144 | -35.8314
25.7661 | -35.5147 | -35.837
25.7722 | -35.515 | -35.8424
25.7783 | -35.5153 | -35.8483
25.7844 | -35.5156 | -35.8536
25.7905 | -35.5158 | -35.8593
25.7966 | -35.5161 | -35.8648
25.8027 | -35.5164 | -35.8705
25.8088 | -35.5167 | -35.8763
25.8149 | -35.517 | -35.8815
25.821 | -35.5173 | -35.8868
25.8271 | -35.5176 | -35.8925
25.8332 | -35.5178 | -35.8979
25.8393 | -35.5181 | -35.9034
25.8454 | -35.5185 | -35.9087
25.8515 | -35.5188 | -35.914
25.8576 | -35.5191 | -35.9196
25.8638 | -35.5193 | -35.9251
25.8699 | -35.5196 | -35.9311
25.876 | -35.5199 | -35.9367
25.8821 | -35.5202 | -35.9427
25.8882 | -35.5205 | -35.9483
25.8943 | -35.5208 | -35.9538
25.9004 | -35.521 | -35.9595
25.9065 | -35.5214 | -35.9649
25.9126 | -35.5216 | -35.9702
25.9187 | -35.5219 | -35.9756
25.9248 | -35.5223 | -35.9813
25.9309 | -35.5225 | -35.987
25.937 | -35.5229 | -35.9926
25.9431 | -35.5231 | -35.9984
25.9492 | -35.5235 | -36.0038
25.9553 | -35.5238 | -36.0095
25.9614 | -35.5241 | -36.0147
25.9675 | -35.5243 | -36.0202
25.9736 | -35.5247 | -36.0259
25.9797 | -35.525 | -36.0316
25.9858 | -35.5252 | -36.037
25.9919 | -35.5255 | -36.0427
25.998 | -35.5259 | -36.0478
26.0042 | -35.5262 | -36.0533
26.0103 | -35.5265 | -36.059
26.0164 | -35.5267 | -36.0641
26.0225 | -35.5271 | -36.0695
26.0286 | -35.5273 | -36.0752
26.0347 | -35.5277 | -36.0806
26.0408 | -35.528 | -36.0864
26.0469 | -35.5283 | -36.0919
26.053 | -35.5286 | -36.0974
26.0591 | -35.5289 | -36.103
26.0652 | -35.5292 | -36.1085
26.0713 | -35.5295 | -36.1143
26.0774 | -35.5299 | -36.12
26.0835 | -35.5302 | -36.1254
26.0896 | -35.5305 | -36.1309
26.0957 | -35.5308 | -36.1364
26.1018 | -35.5311 | -36.1417
26.1079 | -35.5314 | -36.1472
26.114 | -35.5316 | -36.1528
26.1201 | -35.5319 | -36.1586
26.1262 | -35.5322 | -36.1641
26.1323 | -35.5325 | -36.1699
26.1384 | -35.5329 | -36.1752
26.1445 | -35.5332 | -36.1807
26.1507 | -35.5335 | -36.1862
26.1568 | -35.5338 | -36.1918
26.1629 | -35.5341 | -36.1972
26.169 | -35.5343 | -36.2027
26.1751 | -35.5346 | -36.2081
26.1812 | -35.5349 | -36.2137
26.1873 | -35.5352 | -36.2191
26.1934 | -35.5355 | -36.2244
26.1995 | -35.5358 | -36.2297
26.2056 | -35.5361 | -36.2353
26.2117 | -35.5364 | -36.2409
26.2178 | -35.5367 | -36.2464
26.2239 | -35.537 | -36.2518
26.23 | -35.5373 | -36.2576
26.2361 | -35.5376 | -36.2631
26.2422 | -35.5379 | -36.2684
26.2483 | -35.5383 | -36.2736
26.2544 | -35.5386 | -36.279
26.2605 | -35.5389 | -36.2848
26.2666 | -35.5392 | -36.2901
26.2727 | -35.5394 | -36.2956
26.2788 | -35.5397 | -36.3008
26.2849 | -35.54 | -36.3064
26.2911 | -35.5402 | -36.3118
26.2972 | -35.5405 | -36.3174
26.3033 | -35.5408 | -36.3227
26.3094 | -35.5411 | -36.3286
26.3155 | -35.5414 | -36.3339
26.3216 | -35.5417 | -36.3392
26.3277 | -35.542 | -36.3444
26.3338 | -35.5423 | -36.3501
26.3399 | -35.5426 | -36.3551
26.346 | -35.5429 | -36.3603
26.3521 | -35.5432 | -36.3661
26.3582 | -35.5434 | -36.3717
26.3643 | -35.5437 | -36.377
26.3704 | -35.544 | -36.3824
26.3765 | -35.5443 | -36.3879
26.3826 | -35.5447 | -36.3935
26.3887 | -35.545 | -36.399
26.3948 | -35.5453 | -36.4049
26.4009 | -35.5455 | -36.4102
26.407 | -35.5458 | -36.4159
26.4131 | -35.5461 | -36.4213
26.4192 | -35.5464 | -36.4269
26.4253 | -35.5468 | -36.4322
26.4314 | -35.5471 | -36.4374
26.4376 | -35.5474 | -36.4427
26.4437 | -35.5477 | -36.4484
26.4498 | -35.548 | -36.4539
26.4559 | -35.5483 | -36.4591
26.462 | -35.5485 | -36.4643
26.4681 | -35.5488 | -36.4698
26.4742 | -35.5491 | -36.4756
26.4803 | -35.5494 | -36.4813
26.4864 | -35.5498 | -36.4867
26.4925 | -35.55 | -36.492
26.4986 | -35.5503 | -36.4976
26.5047 | -35.5507 | -36.503
26.5108 | -35.551 | -36.5085
26.5169 | -35.5513 | -36.514
26.523 | -35.5517 | -36.5193
26.5291 | -35.552 | -36.5248
26.5352 | -35.5522 | -36.5302
26.5413 | -35.5525 | -36.5357
26.5474 | -35.5529 | -36.5412
26.5535 | -35.5532 | -36.5467
26.5596 | -35.5535 | -36.5519
26.5657 | -35.5538 | -36.5571
26.5718 | -35.5541 | -36.5627
26.578 | -35.5544 | -36.5684
26.5841 | -35.5548 | -36.5741
26.5902 | -35.555 | -36.5794
26.5963 | -35.5553 | -36.5847
26.6024 | -35.5555 | -36.5899
26.6085 | -35.5558 | -36.5957
26.6146 | -35.5561 | -36.6007
26.6207 | -35.5564 | -36.6064
26.6268 | -35.5567 | -36.6119
26.6329 | -35.557 | -36.6175
26.639 | -35.5573 | -36.6232
26.6451 | -35.5576 | -36.6287
26.6512 | -35.5579 | -36.6343
26.6573 | -35.5582 | -36.6399
26.6634 | -35.5585 | -36.6449
26.6695 | -35.5588 | -36.6505
26.6756 | -35.5591 | -36.6561
26.6817 | -35.5594 | -36.6618
26.6878 | -35.5597 | -36.6672
26.6939 | -35.56 | -36.6729
26.7 | -35.5603 | -36.6785
26.7061 | -35.5606 | -36.6839
26.7122 | -35.5609 | -36.6896
26.7183 | -35.5613 | -36.6949
26.7245 | -35.5616 | -36.7005
26.7306 | -35.5619 | -36.706
26.7367 | -35.5622 | -36.7114
26.7428 | -35.5625 | -36.7169
26.7489 | -35.5628 | -36.7223
26.755 | -35.5631 | -36.7272
26.7611 | -35.5634 | -36.7325
26.7672 | -35.5637 | -36.7379
26.7733 | -35.564 | -36.7434
26.7794 | -35.5643 | -36.749
26.7855 | -35.5646 | -36.7547
26.7916 | -35.565 | -36.7597
26.7977 | -35.5653 | -36.7656
26.8038 | -35.5656 | -36.771
26.8099 | -35.5659 | -36.7764
26.816 | -35.5662 | -36.7816
26.8221 | -35.5665 | -36.7868
26.8282 | -35.5669 | -36.7921
26.8343 | -35.5671 | -36.7977
26.8404 | -35.5674 | -36.8031
26.8465 | -35.5677 | -36.8084
26.8526 | -35.568 | -36.8139
26.8587 | -35.5683 | -36.8195
26.8649 | -35.5686 | -36.8247
26.871 | -35.5689 | -36.8303
26.8771 | -35.5692 | -36.8357
26.8832 | -35.5696 | -36.8414
26.8893 | -35.5698 | -36.8469
26.8954 | -35.5701 | -36.8522
26.9015 | -35.5704 | -36.8574
26.9076 | -35.5707 | -36.8627
26.9137 | -35.5711 | -36.8682
26.9198 | -35.5714 | -36.874
26.9259 | -35.5717 | -36.8796
26.932 | -35.5719 | -36.8847
26.9381 | -35.5723 | -36.8906
26.9442 | -35.5726 | -36.8954
26.9503 | -35.5728 | -36.9009
26.9564 | -35.5732 | -36.9066
26.9625 | -35.5735 | -36.912
26.9686 | -35.5738 | -36.9173
26.9747 | -35.5741 | -36.9226
26.9808 | -35.5745 | -36.9278
26.9869 | -35.5748 | -36.9333
26.993 | -35.5751 | -36.9385
26.9991 | -35.5754 | -36.944
27.0052 | -35.5757 | -36.9497
27.0114 | -35.576 | -36.9555
27.0175 | -35.5764 | -36.9607
27.0236 | -35.5767 | -36.966
27.0297 | -35.577 | -36.9715
27.0358 | -35.5773 | -36.9772
27.0419 | -35.5776 | -36.9828
27.048 | -35.5779 | -36.9879
27.0541 | -35.5782 | -36.9931
27.0602 | -35.5785 | -36.9985
27.0663 | -35.5788 | -37.0042
27.0724 | -35.5791 | -37.0094
27.0785 | -35.5795 | -37.0151
27.0846 | -35.5798 | -37.0205
27.0907 | -35.5802 | -37.0258
27.0968 | -35.5805 | -37.0313
27.1029 | -35.5808 | -37.0365
27.109 | -35.5811 | -37.0418
27.1151 | -35.5815 | -37.047
27.1212 | -35.5818 | -37.0519
27.1273 | -35.582 | -37.0574
27.1334 | -35.5823 | -37.0629
27.1395 | -35.5827 | -37.0682
27.1456 | -35.583 | -37.0738
27.1518 | -35.5832 | -37.0793
27.1579 | -35.5836 | -37.0843
27.164 | -35.5839 | -37.0899
27.1701 | -35.5842 | -37.0951
27.1762 | -35.5844 | -37.1002
27.1823 | -35.5847 | -37.1055
27.1884 | -35.585 | -37.1112
27.1945 | -35.5853 | -37.1165
27.2006 | -35.5856 | -37.1218
27.2067 | -35.5859 | -37.1268
27.2128 | -35.5862 | -37.1321
27.2189 | -35.5865 | -37.1378
27.225 | -35.5868 | -37.1432
27.2311 | -35.5871 | -37.1484
27.2372 | -35.5875 | -37.1542
27.2433 | -35.5878 | -37.1596
27.2494 | -35.5882 | -37.1652
27.2555 | -35.5884 | -37.1709
27.2616 | -35.5888 | -37.1762
27.2677 | -35.5891 | -37.1815
27.2738 | -35.5894 | -37.1868
27.2799 | -35.5897 | -37.1922
27.286 | -35.59 | -37.1976
27.2921 | -35.5904 | -37.2031
27.2983 | -35.5907 | -37.2083
27.3044 | -35.591 | -37.2138
27.3105 | -35.5913 | -37.2197
27.3166 | -35.5916 | -37.225
27.3227 | -35.592 | -37.2302
27.3288 | -35.5922 | -37.2354
27.3349 | -35.5926 | -37.2407
27.341 | -35.5929 | -37.2463
27.3471 | -35.5932 | -37.252
27.3532 | -35.5935 | -37.2576
27.3593 | -35.5938 | -37.2632
27.3654 | -35.5942 | -37.2687
27.3715 | -35.5945 | -37.2737
27.3776 | -35.5948 | -37.2795
27.3837 | -35.5952 | -37.2847
27.3898 | -35.5955 | -37.2898
27.3959 | -35.5958 | -37.2953
27.402 | -35.5962 | -37.3005
27.4081 | -35.5965 | -37.3062
27.4142 | -35.5968 | -37.3114
27.4203 | -35.5972 | -37.3166
27.4264 | -35.5974 | -37.322
27.4325 | -35.5978 | -37.3271
27.4387 | -35.5981 | -37.3325
27.4448 | -35.5984 | -37.3383
27.4509 | -35.5987 | -37.3437
27.457 | -35.599 | -37.349
27.4631 | -35.5994 | -37.3546
27.4692 | -35.5997 | -37.3598
27.4753 | -35.6 | -37.3648
27.4814 | -35.6003 | -37.37
27.4875 | -35.6006 | -37.3753
27.4936 | -35.601 | -37.3807
27.4997 | -35.6014 | -37.3863
27.5058 | -35.6017 | -37.3913
27.5119 | -35.6021 | -37.3966
27.518 | -35.6024 | -37.4023
27.5241 | -35.6027 | -37.4079
27.5302 | -35.6031 | -37.4133
27.5363 | -35.6034 | -37.4189
27.5424 | -35.6038 | -37.4241
27.5485 | -35.6041 | -37.4293
27.5546 | -35.6044 | -37.4346
27.5607 | -35.6047 | -37.4402
27.5668 | -35.605 | -37.4453
27.5729 | -35.6053 | -37.4506
27.5791 | -35.6057 | -37.4559
27.5852 | -35.606 | -37.4612
27.5913 | -35.6063 | -37.4666
27.5974 | -35.6066 | -37.4721
27.6035 | -35.6069 | -37.4773
27.6096 | -35.6072 | -37.4823
27.6157 | -35.6075 | -37.4877
27.6218 | -35.6079 | -37.4931
27.6279 | -35.6082 | -37.4989
27.634 | -35.6085 | -37.5044
27.6401 | -35.6088 | -37.5099
27.6462 | -35.6091 | -37.515
27.6523 | -35.6094 | -37.5204
27.6584 | -35.6097 | -37.5254
27.6645 | -35.6101 | -37.5305
27.6706 | -35.6104 | -37.5356
27.6767 | -35.6107 | -37.5409
27.6828 | -35.6111 | -37.5462
27.6889 | -35.6114 | -37.5516
27.695 | -35.6117 | -37.5569
27.7011 | -35.612 | -37.5622
27.7072 | -35.6124 | -37.5675
27.7133 | -35.6127 | -37.573
27.7194 | -35.6131 | -37.5781
27.7256 | -35.6134 | -37.5834
27.7317 | -35.6137 | -37.5885
27.7378 | -35.614 | -37.5938
27.7439 | -35.6144 | -37.5994
27.75 | -35.6147 | -37.605
27.7561 | -35.615 | -37.6102
27.7622 | -35.6153 | -37.6155
27.7683 | -35.6157 | -37.621
27.7744 | -35.616 | -37.6264
27.7805 | -35.6164 | -37.6317
27.7866 | -35.6167 | -37.6372
27.7927 | -35.617 | -37.6424
27.7988 | -35.6173 | -37.6476
27.8049 | -35.6176 | -37.653
27.811 | -35.6179 | -37.6581
27.8171 | -35.6182 | -37.6636
27.8232 | -35.6186 | -37.6689
27.8293 | -35.6189 | -37.6742
27.8354 | -35.6192 | -37.6797
27.8415 | -35.6195 | -37.685
27.8476 | -35.6199 | -37.6901
27.8537 | -35.6202 | -37.6956
27.8598 | -35.6205 | -37.7008
27.866 | -35.6208 | -37.706
27.8721 | -35.6212 | -37.7114
27.8782 | -35.6215 | -37.717
27.8843 | -35.6218 | -37.7225
27.8904 | -35.6222 | -37.7275
27.8965 | -35.6225 | -37.7329
27.9026 | -35.6228 | -37.738
27.9087 | -35.6232 | -37.7432
27.9148 | -35.6235 | -37.7485
27.9209 | -35.6239 | -37.7538
27.927 | -35.6242 | -37.7592
27.9331 | -35.6246 | -37.7647
27.9392 | -35.6249 | -37.7701
27.9453 | -35.6253 | -37.7752
27.9514 | -35.6255 | -37.7805
27.9575 | -35.6258 | -37.7857
27.9636 | -35.6261 | -37.7907
27.9697 | -35.6264 | -37.7959
27.9758 | -35.6268 | -37.8016
27.9819 | -35.6271 | -37.8071
27.988 | -35.6274 | -37.8123
27.9941 | -35.6277 | -37.8178
28.0002 | -35.628 | -37.8229
28.0063 | -35.6283 | -37.8286
28.0125 | -35.6286 | -37.8339
28.0186 | -35.6289 | -37.8395
28.0247 | -35.6293 | -37.8443
28.0308 | -35.6296 | -37.8493
28.0369 | -35.63 | -37.8546
28.043 | -35.6304 | -37.8598
28.0491 | -35.6307 | -37.8652
28.0552 | -35.631 | -37.8702
28.0613 | -35.6314 | -37.8757
28.0674 | -35.6318 | -37.8806
28.0735 | -35.6321 | -37.8857

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.3224 | -42.0202
21.4908 | -35.5008 | -34.8947
16.4948 | -34.8681 | -27.5553
12.6602 | -34.5133 | -21.297
9.7171 | -34.2957 | -16.404
7.4582 | -34.1538 | -12.433
5.7244 | -34.0414 | -9.2091
4.3936 | -33.9812 | -6.6185
3.3722 | -33.9661 | -4.5726
2.5883 | -33.9446 | -3.0225
1.9866 | -33.9354 | -1.7858
1.5248 | -33.9426 | -0.79441
1.1703 | -33.9658 | -0.14786
1 | -33.9575 | 0.17164
0.89822 | -33.9666 | 0.4409
0.68941 | -33.9479 | 1.0211
0.52919 | -33.9433 | 1.4696
0.40612 | -33.9652 | 2.0727
0.31171 | -33.9785 | 2.5247
0.23928 | -33.9906 | 3.2483
0.18358 | -33.9931 | 3.6717
0.14091 | -33.9905 | 4.5532
0.10819 | -34.0188 | 5.7096
0.083017 | -34.0365 | 7.1859
0.063705 | -34.0628 | 9.0562
0.048923 | -34.1285 | 11.5563
0.037479 | -34.2415 | 14.7389
0.028801 | -34.4158 | 18.7502
0.022078 | -34.7228 | 24.0884
0.017023 | -35.1983 | 30.4904
0.013018 | -35.9889 | 38.1876
0.010014 | -37.1177 | 46.0912

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003712 | 1.6842
0.007423 | -33.7401
0.011135 | -37.3874
0.014846 | -40.0496
0.018558 | -43.3781
0.02227 | -47.0104
0.025981 | -48.0668
0.029693 | -49.5016
0.033404 | -50.765
0.037116 | -51.9053
0.040828 | -54.2823
0.044539 | -55.6791
0.048251 | -56.6951
0.051963 | -57.3434
0.055674 | -58.8294
0.059386 | -58.5894
0.063097 | -58.1616
0.066809 | -58.4219
0.070521 | -58.588
0.074232 | -59.6526
0.077944 | -60.6375
0.081655 | -61.1116
0.085367 | -62.1886
0.089079 | -63.0551
0.09279 | -63.423
0.096502 | -63.5578
0.10021 | -63.4184
0.10393 | -64.1949
0.10764 | -65.1373
0.11135 | -64.9606
0.11506 | -65.1814
0.11877 | -65.505
0.12248 | -65.9286
0.1262 | -66.5811
0.12991 | -66.167
0.13362 | -66.7428
0.13733 | -67.6692
0.14104 | -67.6477
0.14475 | -68.0222
0.14846 | -68.3402
0.15218 | -68.4677
0.15589 | -68.5725
0.1596 | -69.8685
0.16331 | -69.2656
0.16702 | -69.3311
0.17073 | -69.9828
0.17445 | -70.1204
0.17816 | -70.3122
0.18187 | -70.5286
0.18558 | -69.705
0.18929 | -70.0939
0.193 | -70.6804
0.19672 | -70.6562
0.20043 | -70.9954
0.20414 | -71.117
0.20785 | -70.9242
0.21156 | -72.0445
0.21527 | -72.2725
0.21899 | -72.2422
0.2227 | -72.1245
0.22641 | -71.2502
0.23012 | -72.3552
0.23383 | -72.5633
0.23754 | -73.1863
0.24125 | -72.5947
0.24497 | -72.7878
0.24868 | -73.4012
0.25239 | -73.5978
0.2561 | -74.3257
0.25981 | -75.2837
0.26352 | -76.2169
0.26724 | -77.1177
0.27095 | -76.0395
0.27466 | -75.7967
0.27837 | -75.4972
0.28208 | -75.4923
0.28579 | -75.8059
0.28951 | -75.4541
0.29322 | -74.6755
0.29693 | -73.9896
0.30064 | -74.915
0.30435 | -76.2199
0.30806 | -76.1809
0.31178 | -76.2118
0.31549 | -76.1429
0.3192 | -75.9444
0.32291 | -75.7025
0.32662 | -75.7648
0.33033 | -76.3561
0.33404 | -76.1138
0.33776 | -76.7275
0.34147 | -76.6666
0.34518 | -77.0578
0.34889 | -77.4846
0.3526 | -77.5513
0.35631 | -76.9877
0.36003 | -77.3169
0.36374 | -77.6012
0.36745 | -78.0875
0.37116 | -78.1361
0.37487 | -78.4603
0.37858 | -77.2696
0.3823 | -76.9977
0.38601 | -78.5623
0.38972 | -78.5652
0.39343 | -78.8619
0.39714 | -78.7382
0.40085 | -78.1206
0.40457 | -79.39
0.40828 | -79.462
0.41199 | -80.1464
0.4157 | -79.7821
0.41941 | -78.993
0.42312 | -78.9801
0.42684 | -78.9276
0.43055 | -78.8568
0.43426 | -78.0254
0.43797 | -78.1723
0.44168 | -79.0928
0.44539 | -78.765
0.4491 | -79.5274
0.45282 | -80.1111
0.45653 | -79.8602
0.46024 | -79.5853
0.46395 | -79.689
0.46766 | -79.7574
0.47137 | -79.5862
0.47509 | -79.2195
0.4788 | -79.5405
0.48251 | -80.1533
0.48622 | -80.5323
0.48993 | -81.5878
0.49364 | -80.3416
0.49735 | -80.8804
0.50107 | -81.4024
0.50478 | -81.7616
0.50849 | -82.7914
0.5122 | -81.8322
0.51591 | -81.0468
0.51963 | -80.8273
0.52334 | -80.5083
0.52705 | -80.9863
0.53076 | -80.9023
0.53447 | -81.293
0.53818 | -81.7111
0.5419 | -81.5753
0.54561 | -80.9591
0.54932 | -80.6946
0.55303 | -80.954
0.55674 | -81.5925
0.56045 | -82.2441
0.56416 | -82.4224
0.56788 | -82.1964
0.57159 | -82.0232
0.5753 | -82.0678
0.57901 | -81.8707
0.58272 | -81.8304
0.58643 | -82.1231
0.59015 | -82.152
0.59386 | -82.0931
0.59757 | -82.3312
0.60128 | -82.4734
0.60499 | -82.6074
0.6087 | -82.6186
0.61242 | -82.5674
0.61613 | -83.1446
0.61984 | -83.8052
0.62355 | -83.9029
0.62726 | -83.7905
0.63097 | -83.3261
0.63469 | -83.0786
0.6384 | -83.1223
0.64211 | -83.032
0.64582 | -82.9566
0.64953 | -83.3897
0.65324 | -84.2076
0.65695 | -84.1802
0.66067 | -83.1694
0.66438 | -82.6263
0.66809 | -82.7268
0.6718 | -83.0699
0.67551 | -83.2118
0.67922 | -83.145
0.68294 | -83.3005
0.68665 | -84.0398
0.69036 | -84.5593
0.69407 | -84.1779
0.69778 | -84.2289
0.70149 | -84.3258
0.70521 | -83.9986
0.70892 | -83.6955
0.71263 | -83.5024
0.71634 | -83.9143
0.72005 | -84.7105
0.72376 | -85.4398
0.72747 | -85.6204
0.73119 | -85.28
0.7349 | -84.8981
0.73861 | -84.6262
0.74232 | -84.2978
0.74603 | -84.0182
0.74974 | -83.9988
0.75346 | -84.0899
0.75717 | -84.1712
0.76088 | -84.3005
0.76459 | -84.6238
0.7683 | -84.6346
0.77201 | -84.4457
0.77573 | -84.7236
0.77944 | -85.3473
0.78315 | -85.5035
0.78686 | -85.3729
0.79057 | -85.2204
0.79428 | -85.0511
0.798 | -85.0928
0.80171 | -85.1512
0.80542 | -85.516
0.80913 | -85.5711
0.81284 | -85.7167
0.81655 | -86.2545
0.82027 | -86.7477
0.82398 | -86.5476
0.82769 | -85.9539
0.8314 | -85.6492
0.83511 | -86.0681
0.83882 | -86.6607
0.84254 | -86.6993
0.84625 | -86.4958
0.84996 | -86.7214
0.85367 | -86.9666
0.85738 | -86.7265
0.86109 | -86.5103
0.86481 | -86.3906
0.86852 | -86.1438
0.87223 | -86.0504
0.87594 | -86.5419
0.87965 | -86.8323
0.88336 | -86.5467
0.88707 | -86.6584
0.89079 | -86.8659
0.8945 | -86.7293
0.89821 | -86.8605
0.90192 | -87.1556
0.90563 | -86.7518
0.90934 | -86.078
0.91306 | -85.7592
0.91677 | -85.6554
0.92048 | -86.1307
0.92419 | -86.9595
0.9279 | -87.1293
0.93161 | -86.691
0.93532 | -86.3947
0.93904 | -86.505
0.94275 | -86.5968
0.94646 | -86.8799
0.95017 | -87.2546
0.95388 | -87.7151
0.95759 | -88.166
0.96131 | -88.8274
0.96502 | -88.8722
0.96873 | -88.0396
0.97244 | -87.1609
0.97615 | -86.799
0.97986 | -86.8917
0.98358 | -87.2051
0.98729 | -87.4095
0.991 | -86.7213
0.99471 | -85.8611
0.99842 | -85.8141
1.0021 | -86.48
1.0058 | -86.9643
1.0096 | -87.2232
1.0133 | -87.4727
1.017 | -87.913
1.0207 | -88.4698
1.0244 | -88.617
1.0281 | -88.1494
1.0318 | -87.5438
1.0355 | -87.1865
1.0393 | -87.2464
1.043 | -87.8315
1.0467 | -88.8005
1.0504 | -89.2049
1.0541 | -88.8469
1.0578 | -88.5702
1.0615 | -88.3168
1.0652 | -87.7654
1.0689 | -88.0146
1.0727 | -88.2442
1.0764 | -88.2219
1.0801 | -87.9395
1.0838 | -87.683
1.0875 | -87.5072
1.0912 | -87.3303
1.0949 | -87.1507
1.0986 | -87.1361
1.1023 | -87.4045
1.1061 | -87.8746
1.1098 | -88.2936
1.1135 | -88.4408
1.1172 | -88.4978
1.1209 | -88.7262
1.1246 | -88.9705
1.1283 | -89.0838
1.132 | -89.2497
1.1358 | -89.4856
1.1395 | -89.553
1.1432 | -89.3462
1.1469 | -88.907
1.1506 | -88.4273
1.1543 | -88.2708
1.158 | -88.5008
1.1617 | -88.741
1.1654 | -88.8429
1.1692 | -88.9933
1.1729 | -89.1481
1.1766 | -89.2067
1.1803 | -89.2435
1.184 | -89.3129
1.1877 | -89.301
1.1914 | -89.1463
1.1951 | -88.8978
1.1988 | -88.6231
1.2026 | -88.3897
1.2063 | -88.2459
1.21 | -88.2678
1.2137 | -88.5732
1.2174 | -89.0706
1.2211 | -89.4115
1.2248 | -89.3395
1.2285 | -89.0202
1.2323 | -88.7558
1.236 | -88.7049
1.2397 | -88.815
1.2434 | -88.9104
1.2471 | -88.8442
1.2508 | -88.6755
1.2545 | -88.671
1.2582 | -89.0112
1.2619 | -89.5721
1.2657 | -90.0362
1.2694 | -90.1827
1.2731 | -90.0496
1.2768 | -89.8127
1.2805 | -89.6384
1.2842 | -89.5558
1.2879 | -89.4529
1.2916 | -89.2996
1.2954 | -89.3089
1.2991 | -89.7282
1.3028 | -90.5018
1.3065 | -91.1675
1.3102 | -91.2363
1.3139 | -90.8255
1.3176 | -90.3843
1.3213 | -90.0225
1.325 | -89.629
1.3288 | -89.2517
1.3325 | -89.1524
1.3362 | -89.4716
1.3399 | -90.0121
1.3436 | -90.4096
1.3473 | -90.4221
1.351 | -90.105
1.3547 | -89.856
1.3584 | -89.9195
1.3622 | -90.1463
1.3659 | -90.2611
1.3696 | -90.2526
1.3733 | -90.4255
1.377 | -90.8933
1.3807 | -91.2108
1.3844 | -90.9425
1.3881 | -90.4833
1.3919 | -90.4027
1.3956 | -90.6667
1.3993 | -90.833
1.403 | -90.7552
1.4067 | -90.6245
1.4104 | -90.5881
1.4141 | -90.6377
1.4178 | -90.6162
1.4215 | -90.5324
1.4253 | -90.6325
1.429 | -91.0005
1.4327 | -91.4062
1.4364 | -91.4644
1.4401 | -91.1529
1.4438 | -90.8548
1.4475 | -90.7703
1.4512 | -90.7824
1.4549 | -90.6912
1.4587 | -90.5855
1.4624 | -90.695
1.4661 | -90.9226
1.4698 | -90.9747
1.4735 | -90.7817
1.4772 | -90.5461
1.4809 | -90.524
1.4846 | -90.8037
1.4884 | -91.2237
1.4921 | -91.515
1.4958 | -91.4948
1.4995 | -91.1948
1.5032 | -90.9101
1.5069 | -90.8861
1.5106 | -91.0021
1.5143 | -91.0356
1.518 | -90.9925
1.5218 | -90.9556
1.5255 | -90.9141
1.5292 | -90.9411
1.5329 | -91.2352
1.5366 | -91.7839
1.5403 | -92.1814
1.544 | -92.1689
1.5477 | -91.9341
1.5515 | -91.7287
1.5552 | -91.6642
1.5589 | -91.6783
1.5626 | -91.5718
1.5663 | -91.2987
1.57 | -91.0457
1.5737 | -90.9816
1.5774 | -91.1132
1.5811 | -91.3885
1.5849 | -91.7558
1.5886 | -92.1484
1.5923 | -92.4267
1.596 | -92.5043
1.5997 | -92.4608
1.6034 | -92.2037
1.6071 | -91.9265
1.6108 | -91.6661
1.6145 | -91.5636
1.6183 | -91.6178
1.622 | -91.7037
1.6257 | -91.7281
1.6294 | -91.7152
1.6331 | -91.7092
1.6368 | -91.7504
1.6405 | -91.8934
1.6442 | -92.1097
1.648 | -92.2915
1.6517 | -92.2842
1.6554 | -92.0026
1.6591 | -91.5859
1.6628 | -91.321
1.6665 | -91.3978
1.6702 | -91.7524
1.6739 | -92.15
1.6776 | -92.3746
1.6814 | -92.3413
1.6851 | -92.1139
1.6888 | -91.8827
1.6925 | -91.8348
1.6962 | -91.9728
1.6999 | -92.1433
1.7036 | -92.2091
1.7073 | -92.1853
1.7111 | -92.1736
1.7148 | -92.2508
1.7185 | -92.401
1.7222 | -92.5251
1.7259 | -92.5515
1.7296 | -92.5188
1.7333 | -92.4991
1.737 | -92.4873
1.7407 | -92.3987
1.7445 | -92.2153
1.7482 | -92.0582
1.7519 | -92.0902
1.7556 | -92.3292
1.7593 | -92.6502
1.763 | -92.8774
1.7667 | -92.9134
1.7704 | -92.7714
1.7741 | -92.5553
1.7779 | -92.4077
1.7816 | -92.3861
1.7853 | -92.4667
1.789 | -92.6079
1.7927 | -92.7817
1.7964 | -92.9463
1.8001 | -92.9821
1.8038 | -92.8531
1.8076 | -92.6628
1.8113 | -92.5444
1.815 | -92.5659
1.8187 | -92.6894
1.8224 | -92.833
1.8261 | -92.914
1.8298 | -92.9389
1.8335 | -92.9694
1.8372 | -92.9892
1.841 | -92.8938
1.8447 | -92.6948
1.8484 | -92.5585
1.8521 | -92.5884
1.8558 | -92.6663
1.8595 | -92.608
1.8632 | -92.4326
1.8669 | -92.3256
1.8706 | -92.3999
1.8744 | -92.6381
1.8781 | -92.9234
1.8818 | -93.1265
1.8855 | -93.1986
1.8892 | -93.2122
1.8929 | -93.2617
1.8966 | -93.3404
1.9003 | -93.3723
1.9041 | -93.3392
1.9078 | -93.3222
1.9115 | -93.3831
1.9152 | -93.4695
1.9189 | -93.4971
1.9226 | -93.4227
1.9263 | -93.2433
1.93 | -93.028
1.9337 | -92.9125
1.9375 | -92.9891
1.9412 | -93.2081
1.9449 | -93.378
1.9486 | -93.3347
1.9523 | -93.1343
1.956 | -93.0019
1.9597 | -93.1106
1.9634 | -93.4304
1.9672 | -93.7978
1.9709 | -94.0655
1.9746 | -94.1784
1.9783 | -94.1879
1.982 | -94.2206
1.9857 | -94.3508
1.9894 | -94.5167
1.9931 | -94.5167
1.9968 | -94.2888
2.0006 | -93.9688
2.0043 | -93.6837
2.008 | -93.4665
2.0117 | -93.2618
2.0154 | -93.037
2.0191 | -92.8499
2.0228 | -92.7699
2.0265 | -92.8133
2.0303 | -92.9222
2.034 | -92.9989
2.0377 | -93.0017
2.0414 | -92.9717
2.0451 | -92.9791
2.0488 | -93.0209
2.0525 | -93.0493
2.0562 | -93.0456
2.0599 | -93.0172
2.0637 | -93.0083
2.0674 | -93.0862
2.0711 | -93.2649
2.0748 | -93.51
2.0785 | -93.7684
2.0822 | -93.9991
2.0859 | -94.1766
2.0896 | -94.2605
2.0933 | -94.1913
2.0971 | -93.9882
2.1008 | -93.7555
2.1045 | -93.5653
2.1082 | -93.4277
2.1119 | -93.3584
2.1156 | -93.3836
2.1193 | -93.49
2.123 | -93.6151
2.1268 | -93.752
2.1305 | -93.9014
2.1342 | -94.0547
2.1379 | -94.156
2.1416 | -94.2642
2.1453 | -94.4194
2.149 | -94.6132
2.1527 | -94.7923
2.1564 | -94.871
2.1602 | -94.8038
2.1639 | -94.6417
2.1676 | -94.4934
2.1713 | -94.4187
2.175 | -94.443
2.1787 | -94.5428
2.1824 | -94.6868
2.1861 | -94.8266
2.1898 | -94.9171
2.1936 | -94.9406
2.1973 | -94.9073
2.201 | -94.8518
2.2047 | -94.753
2.2084 | -94.6283
2.2121 | -94.4969
2.2158 | -94.3949
2.2195 | -94.3084
2.2233 | -94.214
2.227 | -94.1114
2.2307 | -94.0234
2.2344 | -93.9577
2.2381 | -93.9355
2.2418 | -93.9539
2.2455 | -94.0079
2.2492 | -94.071
2.2529 | -94.1562
2.2567 | -94.2754
2.2604 | -94.4579
2.2641 | -94.6819
2.2678 | -94.9177
2.2715 | -95.1269
2.2752 | -95.2469
2.2789 | -95.2037
2.2826 | -94.9854
2.2864 | -94.6753
2.2901 | -94.389
2.2938 | -94.185
2.2975 | -94.0721
2.3012 | -94.0317
2.3049 | -94.0388
2.3086 | -94.076
2.3123 | -94.1407
2.316 | -94.2399
2.3198 | -94.3519
2.3235 | -94.4364
2.3272 | -94.4787
2.3309 | -94.5033
2.3346 | -94.5295
2.3383 | -94.551
2.342 | -94.5522
2.3457 | -94.5519
2.3494 | -94.5812
2.3532 | -94.6359
2.3569 | -94.6747
2.3606 | -94.6658
2.3643 | -94.6292
2.368 | -94.6044
2.3717 | -94.6348
2.3754 | -94.7405
2.3791 | -94.8877
2.3829 | -95.0347
2.3866 | -95.1686
2.3903 | -95.2929
2.394 | -95.4096
2.3977 | -95.4586
2.4014 | -95.3653
2.4051 | -95.1011
2.4088 | -94.7427
2.4125 | -94.3915
2.4163 | -94.1606
2.42 | -94.1208
2.4237 | -94.2499
2.4274 | -94.4618
2.4311 | -94.6425
2.4348 | -94.7183
2.4385 | -94.7217
2.4422 | -94.7461
2.4459 | -94.8361
2.4497 | -94.9764
2.4534 | -95.121
2.4571 | -95.2106
2.4608 | -95.2451
2.4645 | -95.2543
2.4682 | -95.2521
2.4719 | -95.2536
2.4756 | -95.2847
2.4794 | -95.3449
2.4831 | -95.3954
2.4868 | -95.4259
2.4905 | -95.4485
2.4942 | -95.4661
2.4979 | -95.4359
2.5016 | -95.3224
2.5053 | -95.1714
2.509 | -95.0514
2.5128 | -95.0299
2.5165 | -95.1047
2.5202 | -95.2312
2.5239 | -95.3735
2.5276 | -95.5108
2.5313 | -95.632
2.535 | -95.7302
2.5387 | -95.776
2.5425 | -95.7471
2.5462 | -95.6679
2.5499 | -95.6313
2.5536 | -95.7052
2.5573 | -95.9071
2.561 | -96.1701
2.5647 | -96.3794
2.5684 | -96.4401
2.5721 | -96.3446
2.5759 | -96.1515
2.5796 | -95.9584
2.5833 | -95.8459
2.587 | -95.8596
2.5907 | -95.9605
2.5944 | -96.0637
2.5981 | -96.0873
2.6018 | -96.0046
2.6055 | -95.8503
2.6093 | -95.6972
2.613 | -95.6083
2.6167 | -95.6025
2.6204 | -95.6712
2.6241 | -95.7893
2.6278 | -95.8975
2.6315 | -95.9302
2.6352 | -95.8388
2.639 | -95.6526
2.6427 | -95.4427
2.6464 | -95.2851
2.6501 | -95.2204
2.6538 | -95.2582
2.6575 | -95.3696
2.6612 | -95.4937
2.6649 | -95.5795
2.6686 | -95.5822
2.6724 | -95.5124
2.6761 | -95.3772
2.6798 | -95.2873
2.6835 | -95.2315
2.6872 | -95.2448
2.6909 | -95.3385
2.6946 | -95.4948
2.6983 | -95.673
2.7021 | -95.8655
2.7058 | -96.0503
2.7095 | -96.2086
2.7132 | -96.3359
2.7169 | -96.3887
2.7206 | -96.3689
2.7243 | -96.2782
2.728 | -96.1053
2.7317 | -95.9012
2.7355 | -95.7205
2.7392 | -95.6022
2.7429 | -95.5686
2.7466 | -95.6148
2.7503 | -95.7115
2.754 | -95.8246
2.7577 | -95.9118
2.7614 | -95.9318
2.7651 | -95.8959
2.7689 | -95.8359
2.7726 | -95.7989
2.7763 | -95.7997
2.78 | -95.8061
2.7837 | -95.7968
2.7874 | -95.7704
2.7911 | -95.7367
2.7948 | -95.7141
2.7986 | -95.7229
2.8023 | -95.7639
2.806 | -95.8297
2.8097 | -95.8818
2.8134 | -95.8915
2.8171 | -95.8469
2.8208 | -95.7782
2.8245 | -95.7029
2.8282 | -95.6438
2.832 | -95.6106
2.8357 | -95.616
2.8394 | -95.6663
2.8431 | -95.7589
2.8468 | -95.8774
2.8505 | -95.9922
2.8542 | -96.0693
2.8579 | -96.1126
2.8617 | -96.1219
2.8654 | -96.1366
2.8691 | -96.1639
2.8728 | -96.2229
2.8765 | -96.3074
2.8802 | -96.4169
2.8839 | -96.5269
2.8876 | -96.627
2.8913 | -96.6865
2.8951 | -96.6706
2.8988 | -96.5836
2.9025 | -96.4544
2.9062 | -96.3246
2.9099 | -96.2225
2.9136 | -96.1639
2.9173 | -96.1636
2.921 | -96.2025
2.9247 | -96.2706
2.9285 | -96.3542
2.9322 | -96.432
2.9359 | -96.4851
2.9396 | -96.5119
2.9433 | -96.5063
2.947 | -96.4833
2.9507 | -96.4684
2.9544 | -96.4987
2.9582 | -96.5828
2.9619 | -96.6975
2.9656 | -96.7979
2.9693 | -96.8543
2.973 | -96.8359
2.9767 | -96.7636
2.9804 | -96.6782
2.9841 | -96.5822
2.9878 | -96.4814
2.9916 | -96.4091
2.9953 | -96.3922
2.999 | -96.4497
3.0027 | -96.5626
3.0064 | -96.6963
3.0101 | -96.794
3.0138 | -96.8427
3.0175 | -96.8501
3.0212 | -96.8491
3.025 | -96.8633
3.0287 | -96.8933
3.0324 | -96.9479
3.0361 | -97.011
3.0398 | -97.0434
3.0435 | -97.0238
3.0472 | -96.9699
3.0509 | -96.9185
3.0547 | -96.8956
3.0584 | -96.9032
3.0621 | -96.9256
3.0658 | -96.9411
3.0695 | -96.9344
3.0732 | -96.917
3.0769 | -96.9094
3.0806 | -96.9167
3.0843 | -96.9478
3.0881 | -97.0118
3.0918 | -97.1067
3.0955 | -97.2265
3.0992 | -97.3439
3.1029 | -97.4364
3.1066 | -97.4849
3.1103 | -97.4734
3.114 | -97.4157
3.1178 | -97.3207
3.1215 | -97.2106
3.1252 | -97.133
3.1289 | -97.1248
3.1326 | -97.1832
3.1363 | -97.2842
3.14 | -97.3793
3.1437 | -97.4427
3.1474 | -97.4696
3.1512 | -97.4674
3.1549 | -97.4545
3.1586 | -97.4381
3.1623 | -97.4106
3.166 | -97.3534
3.1697 | -97.2523
3.1734 | -97.1274
3.1771 | -96.9955
3.1808 | -96.8927
3.1846 | -96.8247
3.1883 | -96.7868
3.192 | -96.7708
3.1957 | -96.7897
3.1994 | -96.8374
3.2031 | -96.8914
3.2068 | -96.9178
3.2105 | -96.8192
3.2143 | -96.7345
3.218 | -96.6344
3.2217 | -96.5384
3.2254 | -96.4658
3.2291 | -96.4393
3.2328 | -96.4504
3.2365 | -96.4991
3.2402 | -96.5711
3.2439 | -96.6429
3.2477 | -96.692
3.2514 | -96.7096
3.2551 | -96.6927
3.2588 | -96.6588
3.2625 | -96.6484
3.2662 | -96.6804
3.2699 | -96.7524
3.2736 | -96.8467
3.2774 | -96.9509
3.2811 | -97.0474
3.2848 | -97.1306
3.2885 | -97.2056
3.2922 | -97.2684
3.2959 | -97.3215
3.2996 | -97.3492
3.3033 | -97.3464
3.307 | -97.3091
3.3108 | -97.2567
3.3145 | -97.203
3.3182 | -97.167
3.3219 | -97.1755
3.3256 | -97.238
3.3293 | -97.3471
3.333 | -97.4824
3.3367 | -97.615
3.3404 | -97.7391
3.3442 | -97.8385
3.3479 | -97.9104
3.3516 | -97.9454
3.3553 | -97.9443
3.359 | -97.8972
3.3627 | -97.8068
3.3664 | -97.6907
3.3701 | -97.5676
3.3739 | -97.4494
3.3776 | -97.3564
3.3813 | -97.316
3.385 | -97.3362
3.3887 | -97.4171
3.3924 | -97.5528
3.3961 | -97.7062
3.3998 | -97.853
3.4035 | -97.9697
3.4073 | -98.0416
3.411 | -98.0695
3.4147 | -98.0717
3.4184 | -98.0766
3.4221 | -98.0859
3.4258 | -98.1225
3.4295 | -98.1922
3.4332 | -98.2748
3.4369 | -98.3469
3.4407 | -98.3868
3.4444 | -98.3799
3.4481 | -98.3142
3.4518 | -98.2167
3.4555 | -98.1136
3.4592 | -98.0123
3.4629 | -97.9384
3.4666 | -97.8846
3.4704 | -97.8439
3.4741 | -97.815
3.4778 | -97.7973
3.4815 | -97.7836
3.4852 | -97.7768
3.4889 | -97.7768
3.4926 | -97.7854
3.4963 | -97.7967
3.5 | -97.8058
3.5038 | -97.811
3.5075 | -97.783
3.5112 | -97.7282
3.5149 | -97.666
3.5186 | -97.5994
3.5223 | -97.5505
3.526 | -97.543
3.5297 | -97.5804
3.5335 | -97.6547
3.5372 | -97.7567
3.5409 | -97.8641
3.5446 | -97.9527
3.5483 | -98.0164
3.552 | -98.0536
3.5557 | -98.0617
3.5594 | -98.0459
3.5631 | -98.0196
3.5669 | -97.9851
3.5706 | -97.9587
3.5743 | -97.9588
3.578 | -97.9939
3.5817 | -98.0532
3.5854 | -98.1295
3.5891 | -98.2159
3.5928 | -98.2841
3.5965 | -98.33
3.6003 | -98.356
3.604 | -98.3417
3.6077 | -98.2984
3.6114 | -98.2411
3.6151 | -98.1862
3.6188 | -98.1469
3.6225 | -98.1349
3.6262 | -98.1439
3.63 | -98.1532
3.6337 | -98.1444
3.6374 | -98.1139
3.6411 | -98.0813
3.6448 | -98.059
3.6485 | -98.0595
3.6522 | -98.0858
3.6559 | -98.1303
3.6596 | -98.1825
3.6634 | -98.2349
3.6671 | -98.2879
3.6708 | -98.3418
3.6745 | -98.4126
3.6782 | -98.4867
3.6819 | -98.5662
3.6856 | -98.6375
3.6893 | -98.6749
3.693 | -98.6736
3.6968 | -98.6308
3.7005 | -98.5604
3.7042 | -98.4766
3.7079 | -98.3804
3.7116 | -98.2778
3.7153 | -98.1704
3.719 | -98.0529
3.7227 | -97.9529
3.7265 | -97.8929
3.7302 | -97.876
3.7339 | -97.9125
3.7376 | -97.9894
3.7413 | -98.0945
3.745 | -98.1916
3.7487 | -98.284
3.7524 | -98.3688
3.7561 | -98.4475
3.7599 | -98.5278
3.7636 | -98.594
3.7673 | -98.6404
3.771 | -98.6625
3.7747 | -98.6595
3.7784 | -98.6276
3.7821 | -98.5723
3.7858 | -98.5108
3.7896 | -98.4609
3.7933 | -98.4273
3.797 | -98.4149
3.8007 | -98.4187
3.8044 | -98.4485
3.8081 | -98.4889
3.8118 | -98.5465
3.8155 | -98.6152
3.8192 | -98.6898
3.823 | -98.7507
3.8267 | -98.7901
3.8304 | -98.7928
3.8341 | -98.7644
3.8378 | -98.7192
3.8415 | -98.6618
3.8452 | -98.5945
3.8489 | -98.538
3.8526 | -98.4978
3.8564 | -98.467
3.8601 | -98.445
3.8638 | -98.4438
3.8675 | -98.4683
3.8712 | -98.5162
3.8749 | -98.5733
3.8786 | -98.6376
3.8823 | -98.7057
3.8861 | -98.7826
3.8898 | -98.8557
3.8935 | -98.9085
3.8972 | -98.9262
3.9009 | -98.9163
3.9046 | -98.8883
3.9083 | -98.8403
3.912 | -98.7816
3.9157 | -98.7324
3.9195 | -98.7017
3.9232 | -98.6863
3.9269 | -98.6838
3.9306 | -98.6881
3.9343 | -98.7086
3.938 | -98.7449
3.9417 | -98.7897
3.9454 | -98.8452
3.9492 | -98.9082
3.9529 | -98.986
3.9566 | -99.079
3.9603 | -99.1818
3.964 | -99.275
3.9677 | -99.3524
3.9714 | -99.3984
3.9751 | -99.3931
3.9788 | -99.3415
3.9826 | -99.2559
3.9863 | -99.1415
3.99 | -99.0271
3.9937 | -98.9277
3.9974 | -98.8555
4.0011 | -98.8086
4.0048 | -98.7972
4.0085 | -98.8226
4.0122 | -98.8664
4.016 | -98.9212
4.0197 | -98.9739
4.0234 | -99.0035
4.0271 | -99.0211
4.0308 | -99.0351
4.0345 | -99.0477
4.0382 | -99.0646
4.0419 | -99.0828
4.0457 | -99.0903
4.0494 | -99.0883
4.0531 | -99.0731
4.0568 | -99.052
4.0605 | -99.0269
4.0642 | -99.0082
4.0679 | -98.998
4.0716 | -98.9997
4.0753 | -99.0323
4.0791 | -99.092
4.0828 | -99.1778
4.0865 | -99.2874
4.0902 | -99.3974
4.0939 | -99.4978
4.0976 | -99.5869
4.1013 | -99.6522
4.105 | -99.6941
4.1087 | -99.7203
4.1125 | -99.7197
4.1162 | -99.6941
4.1199 | -99.6642
4.1236 | -99.6275
4.1273 | -99.5988
4.131 | -99.5654
4.1347 | -99.5376
4.1384 | -99.5148
4.1422 | -99.487
4.1459 | -99.4544
4.1496 | -99.4179
4.1533 | -99.3736
4.157 | -99.3427
4.1607 | -99.3246
4.1644 | -99.3114
4.1681 | -99.2979
4.1718 | -99.2831
4.1756 | -99.2579
4.1793 | -99.2302
4.183 | -99.2042
4.1867 | -99.1851
4.1904 | -99.1615
4.1941 | -99.1447
4.1978 | -99.1245
4.2015 | -99.0927
4.2053 | -99.0547
4.209 | -99.0179
4.2127 | -98.9928
4.2164 | -98.9872
4.2201 | -98.9961
4.2238 | -99.0306
4.2275 | -99.0869
4.2312 | -99.1541
4.2349 | -99.2152
4.2387 | -99.259
4.2424 | -99.27
4.2461 | -99.2493
4.2498 | -99.1942
4.2535 | -99.1141
4.2572 | -99.0274
4.2609 | -98.9484
4.2646 | -98.8982
4.2683 | -98.8845
4.2721 | -98.8924
4.2758 | -98.9247
4.2795 | -98.9825
4.2832 | -99.0558
4.2869 | -99.1202
4.2906 | -99.1883
4.2943 | -99.2604
4.298 | -99.3303
4.3018 | -99.4005
4.3055 | -99.471
4.3092 | -99.5309
4.3129 | -99.5873
4.3166 | -99.6364
4.3203 | -99.6743
4.324 | -99.7002
4.3277 | -99.72
4.3314 | -99.7276
4.3352 | -99.7279
4.3389 | -99.734
4.3426 | -99.7403
4.3463 | -99.7482
4.35 | -99.7718
4.3537 | -99.8197
4.3574 | -99.892
4.3611 | -99.9855
4.3649 | -100.0867
4.3686 | -100.1949
4.3723 | -100.2969
4.376 | -100.3783
4.3797 | -100.4292
4.3834 | -100.437
4.3871 | -100.4215
4.3908 | -100.3636
4.3945 | -100.2712
4.3983 | -100.1614
4.402 | -100.0373
4.4057 | -99.9097
4.4094 | -99.788
4.4131 | -99.6877
4.4168 | -99.6249
4.4205 | -99.6029
4.4242 | -99.621
4.4279 | -99.6611
4.4317 | -99.7207
4.4354 | -99.7941
4.4391 | -99.8626
4.4428 | -99.9173
4.4465 | -99.9482
4.4502 | -99.9588
4.4539 | -99.9534
4.4576 | -99.924
4.4614 | -99.8839
4.4651 | -99.8301
4.4688 | -99.7696
4.4725 | -99.7146
4.4762 | -99.6692
4.4799 | -99.6344
4.4836 | -99.6137
4.4873 | -99.6172
4.491 | -99.6457
4.4948 | -99.692
4.4985 | -99.7636
4.5022 | -99.8499
4.5059 | -99.946
4.5096 | -100.0469
4.5133 | -100.1495
4.517 | -100.2529
4.5207 | -100.3447
4.5244 | -100.4177
4.5282 | -100.4851
4.5319 | -100.5446
4.5356 | -100.5873
4.5393 | -100.6074
4.543 | -100.6151
4.5467 | -100.6025
4.5504 | -100.5635
4.5541 | -100.5013
4.5579 | -100.4092
4.5616 | -100.3158
4.5653 | -100.2379
4.569 | -100.1742
4.5727 | -100.1275
4.5764 | -100.1007
4.5801 | -100.0982
4.5838 | -100.1052
4.5875 | -100.1162
4.5913 | -100.1222
4.595 | -100.1125
4.5987 | -100.0715
4.6024 | -100.0071
4.6061 | -99.9277
4.6098 | -99.8451
4.6135 | -99.7574
4.6172 | -99.6803
4.621 | -99.6246
4.6247 | -99.588
4.6284 | -99.5758
4.6321 | -99.5807
4.6358 | -99.5967
4.6395 | -99.6269
4.6432 | -99.664
4.6469 | -99.6988
4.6506 | -99.7322
4.6544 | -99.7675
4.6581 | -99.8028
4.6618 | -99.8467
4.6655 | -99.9101
4.6692 | -100.0016
4.6729 | -100.1167
4.6766 | -100.2453
4.6803 | -100.3714
4.684 | -100.4905
4.6878 | -100.5848
4.6915 | -100.6444
4.6952 | -100.6687
4.6989 | -100.6683
4.7026 | -100.6543
4.7063 | -100.6413
4.71 | -100.6304
4.7137 | -100.6284
4.7175 | -100.6451
4.7212 | -100.6764
4.7249 | -100.7132
4.7286 | -100.7452
4.7323 | -100.7662
4.736 | -100.7726
4.7397 | -100.7618
4.7434 | -100.735
4.7471 | -100.7061
4.7509 | -100.6787
4.7546 | -100.6553
4.7583 | -100.641
4.762 | -100.6345
4.7657 | -100.6295
4.7694 | -100.622
4.7731 | -100.6063
4.7768 | -100.5794
4.7806 | -100.5428
4.7843 | -100.493
4.788 | -100.4327
4.7917 | -100.374
4.7954 | -100.3283
4.7991 | -100.2949
4.8028 | -100.2781
4.8065 | -100.2789
4.8102 | -100.2948
4.814 | -100.3254
4.8177 | -100.3628
4.8214 | -100.3839
4.8251 | -100.4032
4.8288 | -100.4192
4.8325 | -100.4176
4.8362 | -100.4028
4.8399 | -100.3796
4.8436 | -100.3511
4.8474 | -100.3194
4.8511 | -100.29
4.8548 | -100.2695
4.8585 | -100.2687
4.8622 | -100.2862
4.8659 | -100.3299
4.8696 | -100.3921
4.8733 | -100.4691
4.8771 | -100.5479
4.8808 | -100.6204
4.8845 | -100.6931
4.8882 | -100.7486
4.8919 | -100.7909
4.8956 | -100.8193
4.8993 | -100.84
4.903 | -100.846
4.9067 | -100.8389
4.9105 | -100.8283
4.9142 | -100.8088
4.9179 | -100.7739
4.9216 | -100.7309
4.9253 | -100.6855
4.929 | -100.6323
4.9327 | -100.574
4.9364 | -100.5156
4.9402 | -100.4696
4.9439 | -100.4411
4.9476 | -100.4426
4.9513 | -100.4739
4.955 | -100.5299
4.9587 | -100.6016
4.9624 | -100.6758
4.9661 | -100.7541
4.9698 | -100.8269
4.9736 | -100.8888
4.9773 | -100.9325
4.981 | -100.968
4.9847 | -100.9874
4.9884 | -100.9922
4.9921 | -100.9778
4.9958 | -100.9515
4.9995 | -100.8991
5.0032 | -100.8334
5.007 | -100.7632
5.0107 | -100.678
5.0144 | -100.5845
5.0181 | -100.492
5.0218 | -100.4078
5.0255 | -100.3378
5.0292 | -100.2921
5.0329 | -100.2718
5.0367 | -100.2712
5.0404 | -100.2888
5.0441 | -100.3293
5.0478 | -100.3701
5.0515 | -100.4127
5.0552 | -100.4459
5.0589 | -100.4545
5.0626 | -100.4488
5.0663 | -100.4303
5.0701 | -100.4079
5.0738 | -100.387
5.0775 | -100.3847
5.0812 | -100.3998
5.0849 | -100.4385
5.0886 | -100.5101
5.0923 | -100.6053
5.096 | -100.7154
5.0997 | -100.8312
5.1035 | -100.9452
5.1072 | -101.0508
5.1109 | -101.1405
5.1146 | -101.2023
5.1183 | -101.2321
5.122 | -101.2237
5.1257 | -101.2001
5.1294 | -101.1567
5.1332 | -101.1113
5.1369 | -101.0728
5.1406 | -101.0443
5.1443 | -101.0361
5.148 | -101.0379
5.1517 | -101.0461
5.1554 | -101.058
5.1591 | -101.0747
5.1628 | -101.0805
5.1666 | -101.0672
5.1703 | -101.0418
5.174 | -101.0024
5.1777 | -100.9591
5.1814 | -100.9106
5.1851 | -100.862
5.1888 | -100.8236
5.1925 | -100.8017
5.1963 | -100.7873
5.2 | -100.7816
5.2037 | -100.7837
5.2074 | -100.797
5.2111 | -100.8279
5.2148 | -100.8687
5.2185 | -100.9132
5.2222 | -100.9588
5.2259 | -101.0095
5.2297 | -101.0637
5.2334 | -101.1085
5.2371 | -101.1403
5.2408 | -101.168
5.2445 | -101.1976
5.2482 | -101.2182
5.2519 | -101.2315
5.2556 | -101.2441
5.2593 | -101.2588
5.2631 | -101.2709
5.2668 | -101.281
5.2705 | -101.2832
5.2742 | -101.2903
5.2779 | -101.2974
5.2816 | -101.2977
5.2853 | -101.2959
5.289 | -101.2826
5.2928 | -101.2664
5.2965 | -101.2512
5.3002 | -101.2375
5.3039 | -101.2313
5.3076 | -101.2292
5.3113 | -101.2423
5.315 | -101.268
5.3187 | -101.314
5.3224 | -101.382
5.3262 | -101.4655
5.3299 | -101.559
5.3336 | -101.6463
5.3373 | -101.7246
5.341 | -101.7917
5.3447 | -101.8436
5.3484 | -101.8776
5.3521 | -101.8502
5.3559 | -101.8385
5.3596 | -101.8087
5.3633 | -101.7701
5.367 | -101.7147
5.3707 | -101.6585
5.3744 | -101.5988
5.3781 | -101.5384
5.3818 | -101.4812
5.3855 | -101.4282
5.3893 | -101.3889
5.393 | -101.3588
5.3967 | -101.337
5.4004 | -101.3197
5.4041 | -101.2991
5.4078 | -101.2864
5.4115 | -101.2838
5.4152 | -101.2886
5.4189 | -101.296
5.4227 | -101.3068
5.4264 | -101.3184
5.4301 | -101.3296
5.4338 | -101.3274
5.4375 | -101.3134
5.4412 | -101.2929
5.4449 | -101.2596
5.4486 | -101.2118
5.4524 | -101.1571
5.4561 | -101.0899
5.4598 | -101.0295
5.4635 | -100.9795
5.4672 | -100.9405
5.4709 | -100.9174
5.4746 | -100.909
5.4783 | -100.9186
5.482 | -100.9446
5.4858 | -100.9876
5.4895 | -101.0451
5.4932 | -101.1064
5.4969 | -101.1714
5.5006 | -101.2362
5.5043 | -101.303
5.508 | -101.3627
5.5117 | -101.4213
5.5155 | -101.4758
5.5192 | -101.5358
5.5229 | -101.5992
5.5266 | -101.6583
5.5303 | -101.7143
5.534 | -101.7632
5.5377 | -101.8072
5.5414 | -101.8438
5.5451 | -101.8665
5.5489 | -101.8699
5.5526 | -101.8669
5.5563 | -101.8563
5.56 | -101.839
5.5637 | -101.8128
5.5674 | -101.7884
5.5711 | -101.7552
5.5748 | -101.7235
5.5785 | -101.6983
5.5823 | -101.6763
5.586 | -101.6623
5.5897 | -101.6478
5.5934 | -101.642
5.5971 | -101.6408
5.6008 | -101.6416
5.6045 | -101.6316
5.6082 | -101.6082
5.612 | -101.583
5.6157 | -101.559
5.6194 | -101.5409
5.6231 | -101.5205
5.6268 | -101.5069
5.6305 | -101.4942
5.6342 | -101.4886
5.6379 | -101.4851
5.6416 | -101.4867
5.6454 | -101.4956
5.6491 | -101.5174
5.6528 | -101.5412
5.6565 | -101.5783
5.6602 | -101.6179
5.6639 | -101.6569
5.6676 | -101.6914
5.6713 | -101.7251
5.675 | -101.7619
5.6788 | -101.7899
5.6825 | -101.8165
5.6862 | -101.842
5.6899 | -101.8604
5.6936 | -101.8653
5.6973 | -101.8557
5.701 | -101.8388
5.7047 | -101.8049
5.7085 | -101.7663
5.7122 | -101.7301
5.7159 | -101.702
5.7196 | -101.6877
5.7233 | -101.6859
5.727 | -101.7034
5.7307 | -101.7429
5.7344 | -101.7951
5.7381 | -101.8643
5.7419 | -101.9405
5.7456 | -102.0146
5.7493 | -102.0921
5.753 | -102.1527
5.7567 | -102.1882
5.7604 | -102.2029
5.7641 | -102.1886
5.7678 | -102.1499
5.7716 | -102.1065
5.7753 | -102.0657
5.779 | -102.0248
5.7827 | -101.981
5.7864 | -101.9488
5.7901 | -101.9371
5.7938 | -101.9394
5.7975 | -101.9458
5.8012 | -101.9672
5.805 | -101.9964
5.8087 | -102.022
5.8124 | -102.044
5.8161 | -102.0523
5.8198 | -102.057
5.8235 | -102.0568
5.8272 | -102.051
5.8309 | -102.0462
5.8346 | -102.0279
5.8384 | -102.0167
5.8421 | -102.0135
5.8458 | -102.004
5.8495 | -101.9918
5.8532 | -101.9692
5.8569 | -101.931
5.8606 | -101.8843
5.8643 | -101.8287
5.8681 | -101.7621
5.8718 | -101.6891
5.8755 | -101.6286
5.8792 | -101.5744
5.8829 | -101.5369
5.8866 | -101.5458
5.8903 | -101.5409
5.894 | -101.5505
5.8977 | -101.5763
5.9015 | -101.6125
5.9052 | -101.668
5.9089 | -101.735
5.9126 | -101.8067
5.9163 | -101.8849
5.92 | -101.9667
5.9237 | -102.0525
5.9274 | -102.1312
5.9311 | -102.1953
5.9349 | -102.2466
5.9386 | -102.2955
5.9423 | -102.3172
5.946 | -102.3081
5.9497 | -102.2818
5.9534 | -102.2331
5.9571 | -102.173
5.9608 | -102.0988
5.9646 | -102.0142
5.9683 | -101.9382
5.972 | -101.8662
5.9757 | -101.7977
5.9794 | -101.7422
5.9831 | -101.6972
5.9868 | -101.6739
5.9905 | -101.6728
5.9942 | -101.6855
5.998 | -101.7107
6.0017 | -101.7589
6.0054 | -101.83
6.0091 | -101.9169
6.0128 | -102.0103
6.0165 | -102.1052
6.0202 | -102.1971
6.0239 | -102.2794
6.0277 | -102.3523
6.0314 | -102.397
6.0351 | -102.4231
6.0388 | -102.4367
6.0425 | -102.4248
6.0462 | -102.4053
6.0499 | -102.3725
6.0536 | -102.326
6.0573 | -102.2778
6.0611 | -102.2319
6.0648 | -102.1862
6.0685 | -102.1538
6.0722 | -102.1301
6.0759 | -102.1098
6.0796 | -102.0914
6.0833 | -102.0823
6.087 | -102.0839
6.0907 | -102.0931
6.0945 | -102.1049
6.0982 | -102.1146
6.1019 | -102.1268
6.1056 | -102.1464
6.1093 | -102.1657
6.113 | -102.1859
6.1167 | -102.2013
6.1204 | -102.2122
6.1242 | -102.2227
6.1279 | -102.2362
6.1316 | -102.2397
6.1353 | -102.2395
6.139 | -102.2273
6.1427 | -102.2065
6.1464 | -102.1911
6.1501 | -102.1746
6.1538 | -102.1574
6.1576 | -102.1431
6.1613 | -102.1235
6.165 | -102.1035
6.1687 | -102.085
6.1724 | -102.0636
6.1761 | -102.0375
6.1798 | -102.0032
6.1835 | -101.9611
6.1873 | -101.9225
6.191 | -101.8986
6.1947 | -101.8831
6.1984 | -101.8728
6.2021 | -101.8669
6.2058 | -101.8678
6.2095 | -101.8806
6.2132 | -101.8946
6.2169 | -101.9101
6.2207 | -101.9214
6.2244 | -101.9279
6.2281 | -101.935
6.2318 | -101.9493
6.2355 | -101.9686
6.2392 | -101.992
6.2429 | -102.0221
6.2466 | -102.0614
6.2503 | -102.1108
6.2541 | -102.1655
6.2578 | -102.2235
6.2615 | -102.2819
6.2652 | -102.3434
6.2689 | -102.4038
6.2726 | -102.4653
6.2763 | -102.5221
6.28 | -102.5666
6.2838 | -102.5951
6.2875 | -102.6171
6.2912 | -102.632
6.2949 | -102.6381
6.2986 | -102.6285
6.3023 | -102.6076
6.306 | -102.5872
6.3097 | -102.5654
6.3134 | -102.5502
6.3172 | -102.5484
6.3209 | -102.5428
6.3246 | -102.5395
6.3283 | -102.5443
6.332 | -102.5479
6.3357 | -102.5517
6.3394 | -102.5493
6.3431 | -102.5428
6.3468 | -102.5298
6.3506 | -102.5103
6.3543 | -102.4872
6.358 | -102.4615
6.3617 | -102.4336
6.3654 | -102.4017
6.3691 | -102.3698
6.3728 | -102.339
6.3765 | -102.3058
6.3803 | -102.2769
6.384 | -102.2511
6.3877 | -102.2306
6.3914 | -102.2161
6.3951 | -102.2017
6.3988 | -102.1846
6.4025 | -102.1728
6.4062 | -102.1618
6.4099 | -102.1486
6.4137 | -102.1458
6.4174 | -102.1444
6.4211 | -102.138
6.4248 | -102.1317
6.4285 | -102.1249
6.4322 | -102.1237
6.4359 | -102.1228
6.4396 | -102.1191
6.4434 | -102.119
6.4471 | -102.1182
6.4508 | -102.1125
6.4545 | -102.1026
6.4582 | -102.099
6.4619 | -102.1018
6.4656 | -102.1109
6.4693 | -102.1254
6.473 | -102.1502
6.4768 | -102.1759
6.4805 | -102.2061
6.4842 | -102.2419
6.4879 | -102.2814
6.4916 | -102.3217
6.4953 | -102.3655
6.499 | -102.4061
6.5027 | -102.4467
6.5064 | -102.486
6.5102 | -102.5187
6.5139 | -102.5527
6.5176 | -102.5884
6.5213 | -102.6243
6.525 | -102.6628
6.5287 | -102.698
6.5324 | -102.732
6.5361 | -102.7708
6.5399 | -102.8114
6.5436 | -102.8485
6.5473 | -102.8865
6.551 | -102.9281
6.5547 | -102.9698
6.5584 | -103.0004
6.5621 | -103.0173
6.5658 | -103.0212
6.5695 | -103.019
6.5733 | -103.011
6.577 | -102.9936
6.5807 | -102.972
6.5844 | -102.9415
6.5881 | -102.9049
6.5918 | -102.8695
6.5955 | -102.8287
6.5992 | -102.7883
6.603 | -102.7524
6.6067 | -102.7189
6.6104 | -102.6891
6.6141 | -102.6603
6.6178 | -102.6364
6.6215 | -102.6225
6.6252 | -102.6083
6.6289 | -102.5958
6.6326 | -102.5868
6.6364 | -102.5795
6.6401 | -102.5751
6.6438 | -102.5745
6.6475 | -102.5745
6.6512 | -102.5797
6.6549 | -102.5897
6.6586 | -102.6029
6.6623 | -102.6295
6.666 | -102.6599
6.6698 | -102.6902
6.6735 | -102.7231
6.6772 | -102.7596
6.6809 | -102.7954
6.6846 | -102.8374
6.6883 | -102.8786
6.692 | -102.9192
6.6957 | -102.9671
6.6995 | -103.011
6.7032 | -103.0461
6.7069 | -103.0755
6.7106 | -103.0949
6.7143 | -103.105
6.718 | -103.0999
6.7217 | -103.0775
6.7254 | -103.0503
6.7291 | -103.0158
6.7329 | -102.9804
6.7366 | -102.9479
6.7403 | -102.9154
6.744 | -102.8851
6.7477 | -102.8621
6.7514 | -102.8446
6.7551 | -102.8323
6.7588 | -102.8248
6.7626 | -102.8153
6.7663 | -102.81
6.77 | -102.8026
6.7737 | -102.7917
6.7774 | -102.783
6.7811 | -102.7702
6.7848 | -102.7487
6.7885 | -102.7259
6.7922 | -102.7042
6.796 | -102.6838
6.7997 | -102.6735
6.8034 | -102.669
6.8071 | -102.6743
6.8108 | -102.6902
6.8145 | -102.7091
6.8182 | -102.7351
6.8219 | -102.757
6.8256 | -102.781
6.8294 | -102.8051
6.8331 | -102.8171
6.8368 | -102.8243
6.8405 | -102.8313
6.8442 | -102.8316
6.8479 | -102.8234
6.8516 | -102.8089
6.8553 | -102.7929
6.8591 | -102.772
6.8628 | -102.7514
6.8665 | -102.7281
6.8702 | -102.7143
6.8739 | -102.7011
6.8776 | -102.6866
6.8813 | -102.6828
6.885 | -102.685
6.8887 | -102.6912
6.8925 | -102.6982
6.8962 | -102.697
6.8999 | -102.6973
6.9036 | -102.7014
6.9073 | -102.6993
6.911 | -102.6867
6.9147 | -102.6757
6.9184 | -102.6681
6.9221 | -102.6642
6.9259 | -102.6548
6.9296 | -102.6483
6.9333 | -102.6504
6.937 | -102.659
6.9407 | -102.6746
6.9444 | -102.7008
6.9481 | -102.7319
6.9518 | -102.766
6.9556 | -102.805
6.9593 | -102.8635
6.963 | -102.9169
6.9667 | -102.9671
6.9704 | -103.0214
6.9741 | -103.0683
6.9778 | -103.1204
6.9815 | -103.1747
6.9852 | -103.2201
6.989 | -103.2543
6.9927 | -103.2814
6.9964 | -103.2988
7.0001 | -103.3018
7.0038 | -103.291
7.0075 | -103.2664
7.0112 | -103.2322
7.0149 | -103.1902
7.0187 | -103.1342
7.0224 | -103.0747
7.0261 | -103.0123
7.0298 | -102.9531
7.0335 | -102.8988
7.0372 | -102.8473
7.0409 | -102.8018
7.0446 | -102.7664
7.0483 | -102.7422
7.0521 | -102.7359
7.0558 | -102.7423
7.0595 | -102.7621
7.0632 | -102.79
7.0669 | -102.8255
7.0706 | -102.8646
7.0743 | -102.9088
7.078 | -102.9524
7.0817 | -102.9979
7.0855 | -103.0362
7.0892 | -103.0615
7.0929 | -103.0765
7.0966 | -103.0898
7.1003 | -103.0899
7.104 | -103.0823
7.1077 | -103.0752
7.1114 | -103.0628
7.1152 | -103.0507
7.1189 | -103.0386
7.1226 | -103.032
7.1263 | -103.0382
7.13 | -103.0432
7.1337 | -103.0544
7.1374 | -103.0764
7.1411 | -103.1012
7.1448 | -103.1303
7.1486 | -103.1607
7.1523 | -103.1924
7.156 | -103.2244
7.1597 | -103.2512
7.1634 | -103.2767
7.1671 | -103.2896
7.1708 | -103.2987
7.1745 | -103.3071
7.1783 | -103.3081
7.182 | -103.3071
7.1857 | -103.298
7.1894 | -103.2795
7.1931 | -103.2618
7.1968 | -103.245
7.2005 | -103.2266
7.2042 | -103.2072
7.2079 | -103.185
7.2117 | -103.1592
7.2154 | -103.134
7.2191 | -103.1081
7.2228 | -103.0895
7.2265 | -103.0782
7.2302 | -103.0657
7.2339 | -103.0518
7.2376 | -103.0471
7.2413 | -103.0486
7.2451 | -103.0455
7.2488 | -103.0456
7.2525 | -103.0464
7.2562 | -103.0491
7.2599 | -103.0479
7.2636 | -103.0454
7.2673 | -103.0478
7.271 | -103.0537
7.2748 | -103.0669
7.2785 | -103.0784
7.2822 | -103.0923
7.2859 | -103.1115
7.2896 | -103.1241
7.2933 | -103.1377
7.297 | -103.1587
7.3007 | -103.1719
7.3044 | -103.1789
7.3082 | -103.1815
7.3119 | -103.1838
7.3156 | -103.187
7.3193 | -103.1877
7.323 | -103.1865
7.3267 | -103.1805
7.3304 | -103.1661
7.3341 | -103.158
7.3378 | -103.1532
7.3416 | -103.1503
7.3453 | -103.1505
7.349 | -103.1526
7.3527 | -103.1605
7.3564 | -103.1779
7.3601 | -103.2038
7.3638 | -103.2275
7.3675 | -103.2552
7.3713 | -103.2834
7.375 | -103.306
7.3787 | -103.3268
7.3824 | -103.3451
7.3861 | -103.3594
7.3898 | -103.3701
7.3935 | -103.3777
7.3972 | -103.3855
7.4009 | -103.3875
7.4047 | -103.3814
7.4084 | -103.3782
7.4121 | -103.3761
7.4158 | -103.3761
7.4195 | -103.3684
7.4232 | -103.3623
7.4269 | -103.3578
7.4306 | -103.3493
7.4344 | -103.3406
7.4381 | -103.3324
7.4418 | -103.3086
7.4455 | -103.2813
7.4492 | -103.2464
7.4529 | -103.2016
7.4566 | -103.1466
7.4603 | -103.0886
7.464 | -103.0221
7.4678 | -102.9523
7.4715 | -102.8896
7.4752 | -102.8328
7.4789 | -102.7814
7.4826 | -102.7364
7.4863 | -102.7017
7.49 | -102.6803
7.4937 | -102.7094
7.4974 | -102.7182
7.5012 | -102.7351
7.5049 | -102.7632
7.5086 | -102.8041
7.5123 | -102.8601
7.516 | -102.9242
7.5197 | -102.9933
7.5234 | -103.0662
7.5271 | -103.1393
7.5309 | -103.2208
7.5346 | -103.3024
7.5383 | -103.3799
7.542 | -103.4577
7.5457 | -103.5314
7.5494 | -103.5978
7.5531 | -103.665
7.5568 | -103.7211
7.5605 | -103.7688
7.5643 | -103.8088
7.568 | -103.846
7.5717 | -103.868
7.5754 | -103.8844
7.5791 | -103.8923
7.5828 | -103.8929
7.5865 | -103.89
7.5902 | -103.8826
7.594 | -103.8728
7.5977 | -103.8576
7.6014 | -103.839
7.6051 | -103.8215
7.6088 | -103.8068
7.6125 | -103.7933
7.6162 | -103.7837
7.6199 | -103.7727
7.6236 | -103.7598
7.6274 | -103.7482
7.6311 | -103.7394
7.6348 | -103.735
7.6385 | -103.7301
7.6422 | -103.727
7.6459 | -103.7165
7.6496 | -103.7013
7.6533 | -103.6845
7.657 | -103.6631
7.6608 | -103.6442
7.6645 | -103.6238
7.6682 | -103.5981
7.6719 | -103.5705
7.6756 | -103.5438
7.6793 | -103.5146
7.683 | -103.4829
7.6867 | -103.4515
7.6905 | -103.4226
7.6942 | -103.3976
7.6979 | -103.3669
7.7016 | -103.334
7.7053 | -103.3119
7.709 | -103.2922
7.7127 | -103.2738
7.7164 | -103.2605
7.7201 | -103.2506
7.7239 | -103.2481
7.7276 | -103.252
7.7313 | -103.2651
7.735 | -103.2875
7.7387 | -103.3205
7.7424 | -103.357
7.7461 | -103.3967
7.7498 | -103.4396
7.7535 | -103.4893
7.7573 | -103.539
7.761 | -103.5872
7.7647 | -103.634
7.7684 | -103.6789
7.7721 | -103.718
7.7758 | -103.743
7.7795 | -103.766
7.7832 | -103.7905
7.787 | -103.8141
7.7907 | -103.8288
7.7944 | -103.844
7.7981 | -103.8546
7.8018 | -103.8637
7.8055 | -103.8673
7.8092 | -103.8633
7.8129 | -103.859
7.8166 | -103.852
7.8204 | -103.8324
7.8241 | -103.814
7.8278 | -103.8002
7.8315 | -103.7792
7.8352 | -103.7493
7.8389 | -103.7167
7.8426 | -103.6824
7.8463 | -103.6432
7.8501 | -103.6011
7.8538 | -103.5473
7.8575 | -103.492
7.8612 | -103.4365
7.8649 | -103.3772
7.8686 | -103.3174
7.8723 | -103.2675
7.876 | -103.2307
7.8797 | -103.2045
7.8835 | -103.1851
7.8872 | -103.1824
7.8909 | -103.194
7.8946 | -103.2138
7.8983 | -103.2454
7.902 | -103.2843
7.9057 | -103.3319
7.9094 | -103.3824
7.9131 | -103.436
7.9169 | -103.4894
7.9206 | -103.5375
7.9243 | -103.5877
7.928 | -103.6313
7.9317 | -103.6625
7.9354 | -103.6919
7.9391 | -103.7133
7.9428 | -103.7286
7.9466 | -103.741
7.9503 | -103.7492
7.954 | -103.7544
7.9577 | -103.7558
7.9614 | -103.7555
7.9651 | -103.7541
7.9688 | -103.7549
7.9725 | -103.7541
7.9762 | -103.7545
7.98 | -103.7524
7.9837 | -103.7537
7.9874 | -103.7512
7.9911 | -103.7479
7.9948 | -103.7519
7.9985 | -103.7617
8.0022 | -103.7685
8.0059 | -103.7863
8.0097 | -103.8106
8.0134 | -103.834
8.0171 | -103.8664
8.0208 | -103.8965
8.0245 | -103.9249
8.0282 | -103.9544
8.0319 | -103.9744
8.0356 | -103.9946
8.0393 | -104.0116
8.0431 | -104.0221
8.0468 | -104.031
8.0505 | -104.0353
8.0542 | -104.0293
8.0579 | -104.0171
8.0616 | -104.0002
8.0653 | -103.9843
8.069 | -103.966
8.0727 | -103.9407
8.0765 | -103.9123
8.0802 | -103.8831
8.0839 | -103.8521
8.0876 | -103.82
8.0913 | -103.784
8.095 | -103.7489
8.0987 | -103.7171
8.1024 | -103.6827
8.1062 | -103.6481
8.1099 | -103.6152
8.1136 | -103.5797
8.1173 | -103.5502
8.121 | -103.5245
8.1247 | -103.4983
8.1284 | -103.4727
8.1321 | -103.4509
8.1358 | -103.4353
8.1396 | -103.4251
8.1433 | -103.4155
8.147 | -103.4081
8.1507 | -103.4086
8.1544 | -103.4113
8.1581 | -103.4198
8.1618 | -103.4326
8.1655 | -103.4504
8.1692 | -103.472
8.173 | -103.4974
8.1767 | -103.5274
8.1804 | -103.559
8.1841 | -103.5914
8.1878 | -103.6251
8.1915 | -103.6582
8.1952 | -103.6925
8.1989 | -103.7246
8.2027 | -103.7532
8.2064 | -103.7806
8.2101 | -103.8019
8.2138 | -103.8202
8.2175 | -103.8333
8.2212 | -103.8416
8.2249 | -103.8443
8.2286 | -103.8485
8.2323 | -103.8527
8.2361 | -103.8542
8.2398 | -103.8528
8.2435 | -103.8553
8.2472 | -103.8587
8.2509 | -103.8613
8.2546 | -103.8669
8.2583 | -103.873
8.262 | -103.8866
8.2658 | -103.9074
8.2695 | -103.9249
8.2732 | -103.9411
8.2769 | -103.9644
8.2806 | -103.9889
8.2843 | -104.0173
8.288 | -104.047
8.2917 | -104.0774
8.2954 | -104.1113
8.2992 | -104.1398
8.3029 | -104.1677
8.3066 | -104.1911
8.3103 | -104.2076
8.314 | -104.2224
8.3177 | -104.2289
8.3214 | -104.2345
8.3251 | -104.2354
8.3288 | -104.2337
8.3326 | -104.2304
8.3363 | -104.2258
8.34 | -104.2215
8.3437 | -104.2176
8.3474 | -104.2148
8.3511 | -104.2085
8.3548 | -104.1998
8.3585 | -104.1899
8.3623 | -104.1754
8.366 | -104.1542
8.3697 | -104.1322
8.3734 | -104.1117
8.3771 | -104.0806
8.3808 | -104.0426
8.3845 | -104.003
8.3882 | -103.9601
8.3919 | -103.9191
8.3957 | -103.8767
8.3994 | -103.8364
8.4031 | -103.8006
8.4068 | -103.765
8.4105 | -103.7344
8.4142 | -103.7107
8.4179 | -103.6908
8.4216 | -103.6826
8.4253 | -103.6834
8.4291 | -103.6906
8.4328 | -103.7078
8.4365 | -103.7267
8.4402 | -103.7481
8.4439 | -103.7695
8.4476 | -103.7951
8.4513 | -103.8223
8.455 | -103.8473
8.4588 | -103.866
8.4625 | -103.8775
8.4662 | -103.8816
8.4699 | -103.8825
8.4736 | -103.8804
8.4773 | -103.8758
8.481 | -103.8727
8.4847 | -103.8696
8.4884 | -103.8652
8.4922 | -103.8586
8.4959 | -103.854
8.4996 | -103.8526
8.5033 | -103.8489
8.507 | -103.8477
8.5107 | -103.8503
8.5144 | -103.8539
8.5181 | -103.8588
8.5219 | -103.8602
8.5256 | -103.8587
8.5293 | -103.8579
8.533 | -103.8548
8.5367 | -103.8525
8.5404 | -103.8541
8.5441 | -103.854
8.5478 | -103.8503
8.5515 | -103.8501
8.5553 | -103.8514
8.559 | -103.8561
8.5627 | -103.8671
8.5664 | -103.8847
8.5701 | -103.9019
8.5738 | -103.9196
8.5775 | -103.9403
8.5812 | -103.9679
8.5849 | -103.9996
8.5887 | -104.0339
8.5924 | -104.0657
8.5961 | -104.0966
8.5998 | -104.1291
8.6035 | -104.163
8.6072 | -104.1955
8.6109 | -104.2289
8.6146 | -104.2608
8.6184 | -104.2909
8.6221 | -104.3186
8.6258 | -104.3466
8.6295 | -104.3694
8.6332 | -104.3841
8.6369 | -104.3951
8.6406 | -104.3989
8.6443 | -104.4005
8.648 | -104.3992
8.6518 | -104.3959
8.6555 | -104.388
8.6592 | -104.3767
8.6629 | -104.3591
8.6666 | -104.3344
8.6703 | -104.3099
8.674 | -104.2858
8.6777 | -104.2617
8.6815 | -104.2347
8.6852 | -104.2083
8.6889 | -104.18
8.6926 | -104.1578
8.6963 | -104.1398
8.7 | -104.1218
8.7037 | -104.1116
8.7074 | -104.1062
8.7111 | -104.1054
8.7149 | -104.1076
8.7186 | -104.1181
8.7223 | -104.1404
8.726 | -104.1621
8.7297 | -104.1834
8.7334 | -104.2069
8.7371 | -104.2322
8.7408 | -104.262
8.7445 | -104.2912
8.7483 | -104.3184
8.752 | -104.3457
8.7557 | -104.3745
8.7594 | -104.4081
8.7631 | -104.4366
8.7668 | -104.4625
8.7705 | -104.4898
8.7742 | -104.5102
8.778 | -104.53
8.7817 | -104.5488
8.7854 | -104.5707
8.7891 | -104.594
8.7928 | -104.6223
8.7965 | -104.6535
8.8002 | -104.6811
8.8039 | -104.7101
8.8076 | -104.7377
8.8114 | -104.7675
8.8151 | -104.795
8.8188 | -104.8159
8.8225 | -104.8316
8.8262 | -104.841
8.8299 | -104.847
8.8336 | -104.8478
8.8373 | -104.84
8.841 | -104.8297
8.8448 | -104.8153
8.8485 | -104.7928
8.8522 | -104.7676
8.8559 | -104.7384
8.8596 | -104.7042
8.8633 | -104.6674
8.867 | -104.6299
8.8707 | -104.5936
8.8745 | -104.5609
8.8782 | -104.5252
8.8819 | -104.4845
8.8856 | -104.4394
8.8893 | -104.3955
8.893 | -104.3594
8.8967 | -104.3248
8.9004 | -104.2904
8.9041 | -104.2577
8.9079 | -104.2229
8.9116 | -104.1897
8.9153 | -104.1563
8.919 | -104.1248
8.9227 | -104.0985
8.9264 | -104.08
8.9301 | -104.0692
8.9338 | -104.0655
8.9376 | -104.071
8.9413 | -104.0818
8.945 | -104.0979
8.9487 | -104.12
8.9524 | -104.1418
8.9561 | -104.1694
8.9598 | -104.2045
8.9635 | -104.2425
8.9672 | -104.2819
8.971 | -104.3157
8.9747 | -104.3483
8.9784 | -104.3775
8.9821 | -104.4005
8.9858 | -104.4198
8.9895 | -104.4359
8.9932 | -104.4449
8.9969 | -104.4512
9.0007 | -104.4529
9.0044 | -104.4522
9.0081 | -104.4552
9.0118 | -104.4586
9.0155 | -104.4656
9.0192 | -104.4644
9.0229 | -104.4655
9.0266 | -104.4683
9.0303 | -104.472
9.0341 | -104.4804
9.0378 | -104.4919
9.0415 | -104.4999
9.0452 | -104.5076
9.0489 | -104.5136
9.0526 | -104.5195
9.0563 | -104.5211
9.06 | -104.5142
9.0637 | -104.4988
9.0675 | -104.4797
9.0712 | -104.4632
9.0749 | -104.4466
9.0786 | -104.4292
9.0823 | -104.4158
9.086 | -104.402
9.0897 | -104.383
9.0934 | -104.3675
9.0972 | -104.3593
9.1009 | -104.3618
9.1046 | -104.3629
9.1083 | -104.3657
9.112 | -104.3732
9.1157 | -104.3834
9.1194 | -104.3922
9.1231 | -104.4001
9.1268 | -104.4132
9.1306 | -104.4276
9.1343 | -104.4367
9.138 | -104.4456
9.1417 | -104.4545
9.1454 | -104.4621
9.1491 | -104.4698
9.1528 | -104.4842
9.1565 | -104.4921
9.1602 | -104.4979
9.164 | -104.5028
9.1677 | -104.4985
9.1714 | -104.4912
9.1751 | -104.4849
9.1788 | -104.4736
9.1825 | -104.4612
9.1862 | -104.4428
9.1899 | -104.4205
9.1937 | -104.4018
9.1974 | -104.3809
9.2011 | -104.3506
9.2048 | -104.3179
9.2085 | -104.2817
9.2122 | -104.2452
9.2159 | -104.2062
9.2196 | -104.1694
9.2233 | -104.1392
9.2271 | -104.1116
9.2308 | -104.0886
9.2345 | -104.0705
9.2382 | -104.0563
9.2419 | -104.0473
9.2456 | -104.0428
9.2493 | -104.0444
9.253 | -104.0529
9.2568 | -104.0725
9.2605 | -104.1001
9.2642 | -104.1303
9.2679 | -104.1678
9.2716 | -104.2134
9.2753 | -104.2626
9.279 | -104.313
9.2827 | -104.3618
9.2864 | -104.4067
9.2902 | -104.452
9.2939 | -104.4946
9.2976 | -104.5351
9.3013 | -104.5704
9.305 | -104.6042
9.3087 | -104.6288
9.3124 | -104.6421
9.3161 | -104.6495
9.3198 | -104.6508
9.3236 | -104.6492
9.3273 | -104.6486
9.331 | -104.6415
9.3347 | -104.636
9.3384 | -104.6309
9.3421 | -104.6168
9.3458 | -104.6065
9.3495 | -104.5935
9.3533 | -104.5797
9.357 | -104.5675
9.3607 | -104.5545
9.3644 | -104.5417
9.3681 | -104.5312
9.3718 | -104.5226
9.3755 | -104.5155
9.3792 | -104.5103
9.3829 | -104.5094
9.3867 | -104.5098
9.3904 | -104.5126
9.3941 | -104.5163
9.3978 | -104.522
9.4015 | -104.5298
9.4052 | -104.5399
9.4089 | -104.5542
9.4126 | -104.5739
9.4163 | -104.5951
9.4201 | -104.6162
9.4238 | -104.6379
9.4275 | -104.6544
9.4312 | -104.6726
9.4349 | -104.6879
9.4386 | -104.6977
9.4423 | -104.7058
9.446 | -104.7152
9.4498 | -104.7233
9.4535 | -104.7256
9.4572 | -104.7234
9.4609 | -104.7172
9.4646 | -104.7063
9.4683 | -104.6874
9.472 | -104.6676
9.4757 | -104.6502
9.4794 | -104.6324
9.4832 | -104.6111
9.4869 | -104.5919
9.4906 | -104.5691
9.4943 | -104.5423
9.498 | -104.5225
9.5017 | -104.5118
9.5054 | -104.5026
9.5091 | -104.4967
9.5129 | -104.4944
9.5166 | -104.4919
9.5203 | -104.4953
9.524 | -104.5064
9.5277 | -104.5217
9.5314 | -104.5429
9.5351 | -104.5646
9.5388 | -104.5889
9.5425 | -104.6168
9.5463 | -104.6483
9.55 | -104.6796
9.5537 | -104.7084
9.5574 | -104.7447
9.5611 | -104.7775
9.5648 | -104.8099
9.5685 | -104.8398
9.5722 | -104.8664
9.5759 | -104.8953
9.5797 | -104.9183
9.5834 | -104.9361
9.5871 | -104.9536
9.5908 | -104.9632
9.5945 | -104.9685
9.5982 | -104.9714
9.6019 | -104.9728
9.6056 | -104.9763
9.6094 | -104.9713
9.6131 | -104.9625
9.6168 | -104.9526
9.6205 | -104.9433
9.6242 | -104.9322
9.6279 | -104.9241
9.6316 | -104.9162
9.6353 | -104.9049
9.639 | -104.8902
9.6428 | -104.8754
9.6465 | -104.8647
9.6502 | -104.8559
9.6539 | -104.85
9.6576 | -104.8487
9.6613 | -104.8439
9.665 | -104.8377
9.6687 | -104.834
9.6725 | -104.8313
9.6762 | -104.8314
9.6799 | -104.8328
9.6836 | -104.8348
9.6873 | -104.8362
9.691 | -104.837
9.6947 | -104.8381
9.6984 | -104.8391
9.7021 | -104.8422
9.7059 | -104.8465
9.7096 | -104.8497
9.7133 | -104.8502
9.717 | -104.8461
9.7207 | -104.8435
9.7244 | -104.8424
9.7281 | -104.8411
9.7318 | -104.8438
9.7355 | -104.8426
9.7393 | -104.8363
9.743 | -104.8275
9.7467 | -104.8236
9.7504 | -104.8211
9.7541 | -104.8188
9.7578 | -104.8157
9.7615 | -104.81
9.7652 | -104.8032
9.769 | -104.8017
9.7727 | -104.8002
9.7764 | -104.7978
9.7801 | -104.799
9.7838 | -104.7986
9.7875 | -104.7912
9.7912 | -104.7882
9.7949 | -104.7862
9.7986 | -104.7819
9.8024 | -104.7753
9.8061 | -104.7692
9.8098 | -104.7639
9.8135 | -104.7605
9.8172 | -104.76
9.8209 | -104.7594
9.8246 | -104.7558
9.8283 | -104.7531
9.8321 | -104.7557
9.8358 | -104.7588
9.8395 | -104.7636
9.8432 | -104.764
9.8469 | -104.7682
9.8506 | -104.7725
9.8543 | -104.7776
9.858 | -104.7843
9.8617 | -104.7907
9.8655 | -104.7957
9.8692 | -104.7976
9.8729 | -104.8012
9.8766 | -104.8049
9.8803 | -104.8078
9.884 | -104.8114
9.8877 | -104.8145
9.8914 | -104.8144
9.8951 | -104.8119
9.8989 | -104.808
9.9026 | -104.8068
9.9063 | -104.8032
9.91 | -104.8011
9.9137 | -104.7982
9.9174 | -104.798
9.9211 | -104.7954
9.9248 | -104.7948
9.9286 | -104.7971
9.9323 | -104.8046
9.936 | -104.8074
9.9397 | -104.8133
9.9434 | -104.8267
9.9471 | -104.8406
9.9508 | -104.8533
9.9545 | -104.8659
9.9582 | -104.8792
9.962 | -104.8897
9.9657 | -104.9028
9.9694 | -104.9211
9.9731 | -104.9336
9.9768 | -104.9414
9.9805 | -104.9455
9.9842 | -104.9519
9.9879 | -104.9613
9.9916 | -104.9634
9.9954 | -104.9614
9.9991 | -104.9597
10.0028 | -104.96
10.0065 | -104.9599
10.0102 | -104.958
10.0139 | -104.9565
10.0176 | -104.9496
10.0213 | -104.9434
10.0251 | -104.9416
10.0288 | -104.9415
10.0325 | -104.9408
10.0362 | -104.9396
10.0399 | -104.9382
10.0436 | -104.9388
10.0473 | -104.9394
10.051 | -104.9402
10.0547 | -104.9404
10.0585 | -104.9416
10.0622 | -104.9444
10.0659 | -104.9502
10.0696 | -104.9522
10.0733 | -104.9572
10.077 | -104.9618
10.0807 | -104.966
10.0844 | -104.9682
10.0882 | -104.9695
10.0919 | -104.9725
10.0956 | -104.9742
10.0993 | -104.9707
10.103 | -104.9652
10.1067 | -104.9613
10.1104 | -104.9563
10.1141 | -104.947
10.1178 | -104.943
10.1216 | -104.9434
10.1253 | -104.9491
10.129 | -104.9471
10.1327 | -104.9471
10.1364 | -104.9545
10.1401 | -104.9676
10.1438 | -104.9798
10.1475 | -104.9906
10.1512 | -105.0029
10.155 | -105.0132
10.1587 | -105.0277
10.1624 | -105.0418
10.1661 | -105.0513
10.1698 | -105.0569
10.1735 | -105.0462
10.1772 | -105.043
10.1809 | -105.0359
10.1847 | -105.0285
10.1884 | -105.0204
10.1921 | -105.0082
10.1958 | -104.9954
10.1995 | -104.9817
10.2032 | -104.9655
10.2069 | -104.9443
10.2106 | -104.9232
10.2143 | -104.9059
10.2181 | -104.8924
10.2218 | -104.8803
10.2255 | -104.8686
10.2292 | -104.8631
10.2329 | -104.8605
10.2366 | -104.8609
10.2403 | -104.8654
10.244 | -104.8747
10.2477 | -104.8867
10.2515 | -104.9004
10.2552 | -104.9181
10.2589 | -104.9371
10.2626 | -104.9596
10.2663 | -104.98
10.27 | -104.9991
10.2737 | -105.0191
10.2774 | -105.0439
10.2812 | -105.0639
10.2849 | -105.0776
10.2886 | -105.0888
10.2923 | -105.1053
10.296 | -105.1185
10.2997 | -105.1228
10.3034 | -105.1285
10.3071 | -105.1303
10.3108 | -105.127
10.3146 | -105.1205
10.3183 | -105.1081
10.322 | -105.0906
10.3257 | -105.0702
10.3294 | -105.0465
10.3331 | -105.0223
10.3368 | -104.9951
10.3405 | -104.9705
10.3443 | -104.9462
10.348 | -104.9203
10.3517 | -104.8944
10.3554 | -104.8704
10.3591 | -104.8459
10.3628 | -104.8266
10.3665 | -104.8097
10.3702 | -104.7946
10.3739 | -104.783
10.3777 | -104.7778
10.3814 | -104.7741
10.3851 | -104.777
10.3888 | -104.7827
10.3925 | -104.7882
10.3962 | -104.7993
10.3999 | -104.8158
10.4036 | -104.8313
10.4073 | -104.8462
10.4111 | -104.8636
10.4148 | -104.8843
10.4185 | -104.9064
10.4222 | -104.9283
10.4259 | -104.9497
10.4296 | -104.9636
10.4333 | -104.9792
10.437 | -104.9959
10.4408 | -105.0054
10.4445 | -105.0081
10.4482 | -105.0063
10.4519 | -105.0048
10.4556 | -105.0008
10.4593 | -104.9903
10.463 | -104.9798
10.4667 | -104.9677
10.4704 | -104.9518
10.4742 | -104.9374
10.4779 | -104.9247
10.4816 | -104.9094
10.4853 | -104.8973
10.489 | -104.8841
10.4927 | -104.8684
10.4964 | -104.8586
10.5001 | -104.854
10.5039 | -104.851
10.5076 | -104.85
10.5113 | -104.8554
10.515 | -104.8556
10.5187 | -104.853
10.5224 | -104.8497
10.5261 | -104.8505
10.5298 | -104.8538
10.5335 | -104.8609
10.5373 | -104.8695
10.541 | -104.8791
10.5447 | -104.8944
10.5484 | -104.9086
10.5521 | -104.9216
10.5558 | -104.9379
10.5595 | -104.9525
10.5632 | -104.9685
10.5669 | -104.9841
10.5707 | -104.9992
10.5744 | -105.012
10.5781 | -105.0214
10.5818 | -105.0346
10.5855 | -105.0468
10.5892 | -105.053
10.5929 | -105.0619
10.5966 | -105.0688
10.6004 | -105.0774
10.6041 | -105.0859
10.6078 | -105.089
10.6115 | -105.093
10.6152 | -105.1033
10.6189 | -105.1113
10.6226 | -105.1188
10.6263 | -105.1264
10.63 | -105.1318
10.6338 | -105.1367
10.6375 | -105.1393
10.6412 | -105.1447
10.6449 | -105.1467
10.6486 | -105.1449
10.6523 | -105.142
10.656 | -105.143
10.6597 | -105.1409
10.6634 | -105.137
10.6672 | -105.136
10.6709 | -105.1339
10.6746 | -105.1319
10.6783 | -105.1263
10.682 | -105.1174
10.6857 | -105.1094
10.6894 | -105.0991
10.6931 | -105.0863
10.6969 | -105.0741
10.7006 | -105.0613
10.7043 | -105.0472
10.708 | -105.0346
10.7117 | -105.0214
10.7154 | -105.0126
10.7191 | -105.0009
10.7228 | -104.9882
10.7265 | -104.9825
10.7303 | -104.9803
10.734 | -104.9806
10.7377 | -104.9862
10.7414 | -104.9935
10.7451 | -105.0045
10.7488 | -105.0179
10.7525 | -105.0336
10.7562 | -105.0507
10.76 | -105.0643
10.7637 | -105.0841
10.7674 | -105.1033
10.7711 | -105.1158
10.7748 | -105.1338
10.7785 | -105.1484
10.7822 | -105.1601
10.7859 | -105.167
10.7896 | -105.1707
10.7934 | -105.1693
10.7971 | -105.1617
10.8008 | -105.1553
10.8045 | -105.1415
10.8082 | -105.1222
10.8119 | -105.0997
10.8156 | -105.0777
10.8193 | -105.052
10.823 | -105.0229
10.8268 | -104.9946
10.8305 | -104.9635
10.8342 | -104.9322
10.8379 | -104.9056
10.8416 | -104.8794
10.8453 | -104.8561
10.849 | -104.8379
10.8527 | -104.8224
10.8565 | -104.8111
10.8602 | -104.8043
10.8639 | -104.8027
10.8676 | -104.806
10.8713 | -104.8153
10.875 | -104.8287
10.8787 | -104.8476
10.8824 | -104.8709
10.8861 | -104.8932
10.8899 | -104.92
10.8936 | -104.9516
10.8973 | -104.983
10.901 | -105.0169
10.9047 | -105.0572
10.9084 | -105.0906
10.9121 | -105.124
10.9158 | -105.1602
10.9196 | -105.1979
10.9233 | -105.2228
10.927 | -105.2486
10.9307 | -105.2747
10.9344 | -105.2909
10.9381 | -105.3042
10.9418 | -105.3119
10.9455 | -105.3153
10.9492 | -105.3124
10.953 | -105.3069
10.9567 | -105.2971
10.9604 | -105.2782
10.9641 | -105.2613
10.9678 | -105.2374
10.9715 | -105.2119
10.9752 | -105.1856
10.9789 | -105.1618
10.9826 | -105.1392
10.9864 | -105.1199
10.9901 | -105.1019
10.9938 | -105.0878
10.9975 | -105.0771
11.0012 | -105.0691
11.0049 | -105.0668
11.0086 | -105.0716
11.0123 | -105.0808
11.0161 | -105.0948
11.0198 | -105.1109
11.0235 | -105.124
11.0272 | -105.1442
11.0309 | -105.1661
11.0346 | -105.1892
11.0383 | -105.216
11.042 | -105.2425
11.0457 | -105.2667
11.0495 | -105.2882
11.0532 | -105.3083
11.0569 | -105.3264
11.0606 | -105.3421
11.0643 | -105.3513
11.068 | -105.3677
11.0717 | -105.3828
11.0754 | -105.3897
11.0792 | -105.394
11.0829 | -105.4015
11.0866 | -105.4078
11.0903 | -105.4121
11.094 | -105.4163
11.0977 | -105.4183
11.1014 | -105.4182
11.1051 | -105.4192
11.1088 | -105.4203
11.1126 | -105.4135
11.1163 | -105.4055
11.12 | -105.397
11.1237 | -105.3806
11.1274 | -105.3639
11.1311 | -105.3423
11.1348 | -105.3144
11.1385 | -105.2881
11.1422 | -105.2634
11.146 | -105.2331
11.1497 | -105.199
11.1534 | -105.1598
11.1571 | -105.1221
11.1608 | -105.0847
11.1645 | -105.048
11.1682 | -105.01
11.1719 | -104.9759
11.1757 | -104.9426
11.1794 | -104.9141
11.1831 | -104.8956
11.1868 | -104.8791
11.1905 | -104.8701
11.1942 | -104.8679
11.1979 | -104.8729
11.2016 | -104.8873
11.2053 | -104.9101
11.2091 | -104.9432
11.2128 | -104.9818
11.2165 | -105.0277
11.2202 | -105.0805
11.2239 | -105.1375
11.2276 | -105.1919
11.2313 | -105.2486
11.235 | -105.3102
11.2387 | -105.3691
11.2425 | -105.4088
11.2462 | -105.4605
11.2499 | -105.5027
11.2536 | -105.5359
11.2573 | -105.5653
11.261 | -105.59
11.2647 | -105.6086
11.2684 | -105.6166
11.2722 | -105.6211
11.2759 | -105.6227
11.2796 | -105.6178
11.2833 | -105.6104
11.287 | -105.5966
11.2907 | -105.5749
11.2944 | -105.5477
11.2981 | -105.5247
11.3018 | -105.4911
11.3056 | -105.4537
11.3093 | -105.4162
11.313 | -105.381
11.3167 | -105.349
11.3204 | -105.3149
11.3241 | -105.2809
11.3278 | -105.246
11.3315 | -105.2179
11.3353 | -105.1901
11.339 | -105.1659
11.3427 | -105.1448
11.3464 | -105.1255
11.3501 | -105.1115
11.3538 | -105.0981
11.3575 | -105.0882
11.3612 | -105.079
11.3649 | -105.0735
11.3687 | -105.0719
11.3724 | -105.0694
11.3761 | -105.0708
11.3798 | -105.0763
11.3835 | -105.0835
11.3872 | -105.093
11.3909 | -105.1069
11.3946 | -105.1165
11.3983 | -105.1242
11.4021 | -105.1361
11.4058 | -105.1477
11.4095 | -105.1612
11.4132 | -105.1731
11.4169 | -105.1857
11.4206 | -105.1996
11.4243 | -105.2119
11.428 | -105.2212
11.4318 | -105.2299
11.4355 | -105.2384
11.4392 | -105.2446
11.4429 | -105.2481
11.4466 | -105.251
11.4503 | -105.2501
11.454 | -105.2493
11.4577 | -105.2503
11.4614 | -105.2509
11.4652 | -105.2447
11.4689 | -105.242
11.4726 | -105.239
11.4763 | -105.2367
11.48 | -105.2317
11.4837 | -105.2265
11.4874 | -105.2201
11.4911 | -105.2137
11.4948 | -105.2099
11.4986 | -105.2054
11.5023 | -105.1983
11.506 | -105.1909
11.5097 | -105.1852
11.5134 | -105.1783
11.5171 | -105.1697
11.5208 | -105.1651
11.5245 | -105.1585
11.5283 | -105.1552
11.532 | -105.1528
11.5357 | -105.1498
11.5394 | -105.1484
11.5431 | -105.1447
11.5468 | -105.1396
11.5505 | -105.136
11.5542 | -105.135
11.5579 | -105.1328
11.5617 | -105.1362
11.5654 | -105.1408
11.5691 | -105.1434
11.5728 | -105.1477
11.5765 | -105.1523
11.5802 | -105.1594
11.5839 | -105.1698
11.5876 | -105.1824
11.5914 | -105.1925
11.5951 | -105.2084
11.5988 | -105.2274
11.6025 | -105.2498
11.6062 | -105.2729
11.6099 | -105.3006
11.6136 | -105.3289
11.6173 | -105.3617
11.621 | -105.3961
11.6248 | -105.4268
11.6285 | -105.4568
11.6322 | -105.4877
11.6359 | -105.5149
11.6396 | -105.5409
11.6433 | -105.5584
11.647 | -105.576
11.6507 | -105.589
11.6544 | -105.5991
11.6582 | -105.6076
11.6619 | -105.6104
11.6656 | -105.6086
11.6693 | -105.6063
11.673 | -105.6032
11.6767 | -105.5948
11.6804 | -105.5831
11.6841 | -105.5663
11.6879 | -105.5481
11.6916 | -105.5349
11.6953 | -105.5156
11.699 | -105.4912
11.7027 | -105.4703
11.7064 | -105.448
11.7101 | -105.4272
11.7138 | -105.4069
11.7175 | -105.3851
11.7213 | -105.3639
11.725 | -105.344
11.7287 | -105.3246
11.7324 | -105.311
11.7361 | -105.3003
11.7398 | -105.2876
11.7435 | -105.2794
11.7472 | -105.2749
11.751 | -105.2732
11.7547 | -105.2712
11.7584 | -105.2756
11.7621 | -105.2776
11.7658 | -105.2784
11.7695 | -105.28
11.7732 | -105.28
11.7769 | -105.286
11.7806 | -105.2905
11.7844 | -105.2958
11.7881 | -105.2997
11.7918 | -105.3062
11.7955 | -105.3131
11.7992 | -105.3174
11.8029 | -105.3222
11.8066 | -105.3281
11.8103 | -105.3336
11.814 | -105.3429
11.8178 | -105.3477
11.8215 | -105.3518
11.8252 | -105.3555
11.8289 | -105.3577
11.8326 | -105.3595
11.8363 | -105.3625
11.84 | -105.3706
11.8437 | -105.3782
11.8475 | -105.3844
11.8512 | -105.3904
11.8549 | -105.3982
11.8586 | -105.4075
11.8623 | -105.4146
11.866 | -105.4244
11.8697 | -105.437
11.8734 | -105.4487
11.8771 | -105.4612
11.8809 | -105.4733
11.8846 | -105.4864
11.8883 | -105.5008
11.892 | -105.5146
11.8957 | -105.5272
11.8994 | -105.5356
11.9031 | -105.5389
11.9068 | -105.5421
11.9106 | -105.5487
11.9143 | -105.5508
11.918 | -105.5529
11.9217 | -105.5549
11.9254 | -105.5572
11.9291 | -105.5541
11.9328 | -105.5517
11.9365 | -105.5489
11.9402 | -105.5464
11.944 | -105.5456
11.9477 | -105.5384
11.9514 | -105.5311
11.9551 | -105.5255
11.9588 | -105.5181
11.9625 | -105.5125
11.9662 | -105.5098
11.9699 | -105.5085
11.9736 | -105.5084
11.9774 | -105.5089
11.9811 | -105.5102
11.9848 | -105.5115
11.9885 | -105.5147
11.9922 | -105.5206
11.9959 | -105.5252
11.9996 | -105.5282
12.0033 | -105.5319
12.0071 | -105.5339
12.0108 | -105.5345
12.0145 | -105.5304
12.0182 | -105.5285
12.0219 | -105.5301
12.0256 | -105.5298
12.0293 | -105.5288
12.033 | -105.5228
12.0367 | -105.5148
12.0405 | -105.51
12.0442 | -105.5021
12.0479 | -105.492
12.0516 | -105.4834
12.0553 | -105.4767
12.059 | -105.4686
12.0627 | -105.4587
12.0664 | -105.4493
12.0701 | -105.4421
12.0739 | -105.4358
12.0776 | -105.4296
12.0813 | -105.4244
12.085 | -105.4185
12.0887 | -105.4167
12.0924 | -105.4129
12.0961 | -105.4109
12.0998 | -105.4106
12.1036 | -105.4139
12.1073 | -105.4146
12.111 | -105.4133
12.1147 | -105.4141
12.1184 | -105.418
12.1221 | -105.4225
12.1258 | -105.4217
12.1295 | -105.4162
12.1332 | -105.4121
12.137 | -105.4045
12.1407 | -105.3927
12.1444 | -105.3818
12.1481 | -105.3671
12.1518 | -105.3514
12.1555 | -105.3337
12.1592 | -105.3126
12.1629 | -105.2901
12.1667 | -105.267
12.1704 | -105.2445
12.1741 | -105.2253
12.1778 | -105.2096
12.1815 | -105.1964
12.1852 | -105.1832
12.1889 | -105.1703
12.1926 | -105.1603
12.1963 | -105.1498
12.2001 | -105.1423
12.2038 | -105.1394
12.2075 | -105.1426
12.2112 | -105.1473
12.2149 | -105.1519
12.2186 | -105.1597
12.2223 | -105.1699
12.226 | -105.1798
12.2297 | -105.1896
12.2335 | -105.2045
12.2372 | -105.2231
12.2409 | -105.2412
12.2446 | -105.2573
12.2483 | -105.2729
12.252 | -105.2862
12.2557 | -105.2969
12.2594 | -105.3082
12.2632 | -105.3191
12.2669 | -105.3317
12.2706 | -105.3446
12.2743 | -105.3539
12.278 | -105.362
12.2817 | -105.3762
12.2854 | -105.3896
12.2891 | -105.3989
12.2928 | -105.4051
12.2966 | -105.4133
12.3003 | -105.4233
12.304 | -105.4343
12.3077 | -105.4441
12.3114 | -105.4523
12.3151 | -105.453
12.3188 | -105.4539
12.3225 | -105.4491
12.3262 | -105.4469
12.33 | -105.4438
12.3337 | -105.4407
12.3374 | -105.4387
12.3411 | -105.4325
12.3448 | -105.4247
12.3485 | -105.4162
12.3522 | -105.4058
12.3559 | -105.3957
12.3597 | -105.3864
12.3634 | -105.3764
12.3671 | -105.3715
12.3708 | -105.3663
12.3745 | -105.3636
12.3782 | -105.3609
12.3819 | -105.3591
12.3856 | -105.358
12.3893 | -105.3592
12.3931 | -105.3624
12.3968 | -105.3698
12.4005 | -105.3823
12.4042 | -105.3955
12.4079 | -105.4116
12.4116 | -105.4285
12.4153 | -105.4483
12.419 | -105.4684
12.4228 | -105.4904
12.4265 | -105.5166
12.4302 | -105.5435
12.4339 | -105.5696
12.4376 | -105.5931
12.4413 | -105.6115
12.445 | -105.6253
12.4487 | -105.6393
12.4524 | -105.6541
12.4562 | -105.6679
12.4599 | -105.6742
12.4636 | -105.6795
12.4673 | -105.6838
12.471 | -105.6889
12.4747 | -105.6875
12.4784 | -105.6808
12.4821 | -105.6763
12.4858 | -105.6694
12.4896 | -105.6618
12.4933 | -105.6522
12.497 | -105.6407
12.5007 | -105.6306
12.5044 | -105.6206
12.5081 | -105.601
12.5118 | -105.5787
12.5155 | -105.5537
12.5193 | -105.5288
12.523 | -105.5075
12.5267 | -105.4856
12.5304 | -105.4605
12.5341 | -105.4355
12.5378 | -105.4088
12.5415 | -105.3844
12.5452 | -105.3621
12.5489 | -105.3377
12.5527 | -105.3166
12.5564 | -105.2971
12.5601 | -105.2824
12.5638 | -105.2697
12.5675 | -105.2603
12.5712 | -105.2528
12.5749 | -105.2484
12.5786 | -105.2456
12.5824 | -105.2382
12.5861 | -105.2365
12.5898 | -105.239
12.5935 | -105.2423
12.5972 | -105.2475
12.6009 | -105.2558
12.6046 | -105.2642
12.6083 | -105.272
12.612 | -105.2821
12.6158 | -105.2919
12.6195 | -105.3015
12.6232 | -105.3117
12.6269 | -105.3218
12.6306 | -105.3292
12.6343 | -105.3354
12.638 | -105.3357
12.6417 | -105.3345
12.6454 | -105.3367
12.6492 | -105.3372
12.6529 | -105.3354
12.6566 | -105.3363
12.6603 | -105.3385
12.664 | -105.3359
12.6677 | -105.3342
12.6714 | -105.3343
12.6751 | -105.3352
12.6789 | -105.339
12.6826 | -105.3431
12.6863 | -105.3443
12.69 | -105.3472
12.6937 | -105.3513
12.6974 | -105.3545
12.7011 | -105.3585
12.7048 | -105.3618
12.7085 | -105.3643
12.7123 | -105.3688
12.716 | -105.3713
12.7197 | -105.3709
12.7234 | -105.3678
12.7271 | -105.3653
12.7308 | -105.3655
12.7345 | -105.3642
12.7382 | -105.3597
12.7419 | -105.3509
12.7457 | -105.3431
12.7494 | -105.3357
12.7531 | -105.331
12.7568 | -105.3269
12.7605 | -105.3221
12.7642 | -105.319
12.7679 | -105.3163
12.7716 | -105.312
12.7754 | -105.3097
12.7791 | -105.3095
12.7828 | -105.3126
12.7865 | -105.3176
12.7902 | -105.3234
12.7939 | -105.3316
12.7976 | -105.3417
12.8013 | -105.3528
12.805 | -105.369
12.8088 | -105.3872
12.8125 | -105.4036
12.8162 | -105.4209
12.8199 | -105.4401
12.8236 | -105.4572
12.8273 | -105.4696
12.831 | -105.482
12.8347 | -105.4951
12.8385 | -105.5098
12.8422 | -105.5234
12.8459 | -105.5336
12.8496 | -105.5316
12.8533 | -105.5357
12.857 | -105.539
12.8607 | -105.5424
12.8644 | -105.5426
12.8681 | -105.5424
12.8719 | -105.5458
12.8756 | -105.545
12.8793 | -105.5416
12.883 | -105.5423
12.8867 | -105.5394
12.8904 | -105.5345
12.8941 | -105.5299
12.8978 | -105.5269
12.9016 | -105.5247
12.9053 | -105.5212
12.909 | -105.5171
12.9127 | -105.5156
12.9164 | -105.5137
12.9201 | -105.5107
12.9238 | -105.51
12.9275 | -105.5125
12.9312 | -105.5134
12.935 | -105.5161
12.9387 | -105.5205
12.9424 | -105.523
12.9461 | -105.5284
12.9498 | -105.5354
12.9535 | -105.5372
12.9572 | -105.5398
12.9609 | -105.5405
12.9646 | -105.5357
12.9684 | -105.5294
12.9721 | -105.5248
12.9758 | -105.5185
12.9795 | -105.5099
12.9832 | -105.5015
12.9869 | -105.4906
12.9906 | -105.4755
12.9943 | -105.4628
12.9981 | -105.4541
13.0018 | -105.4465
13.0055 | -105.4399
13.0092 | -105.4339
13.0129 | -105.428
13.0166 | -105.4225
13.0203 | -105.4164
13.024 | -105.4109
13.0277 | -105.4069
13.0315 | -105.4048
13.0352 | -105.4052
13.0389 | -105.4101
13.0426 | -105.4152
13.0463 | -105.4191
13.05 | -105.4231
13.0537 | -105.4295
13.0574 | -105.4373
13.0611 | -105.4485
13.0649 | -105.4649
13.0686 | -105.4802
13.0723 | -105.493
13.076 | -105.507
13.0797 | -105.525
13.0834 | -105.5407
13.0871 | -105.5549
13.0908 | -105.5713
13.0946 | -105.5891
13.0983 | -105.6033
13.102 | -105.6123
13.1057 | -105.6162
13.1094 | -105.6183
13.1131 | -105.6207
13.1168 | -105.6196
13.1205 | -105.6139
13.1242 | -105.607
13.128 | -105.6013
13.1317 | -105.5939
13.1354 | -105.5808
13.1391 | -105.568
13.1428 | -105.5571
13.1465 | -105.5443
13.1502 | -105.5324
13.1539 | -105.5221
13.1577 | -105.5109
13.1614 | -105.4967
13.1651 | -105.4835
13.1688 | -105.473
13.1725 | -105.4612
13.1762 | -105.4524
13.1799 | -105.4484
13.1836 | -105.4441
13.1873 | -105.4365
13.1911 | -105.4322
13.1948 | -105.4318
13.1985 | -105.4337
13.2022 | -105.4369
13.2059 | -105.4408
13.2096 | -105.4468
13.2133 | -105.4549
13.217 | -105.4645
13.2207 | -105.4764
13.2245 | -105.4853
13.2282 | -105.4945
13.2319 | -105.5042
13.2356 | -105.5113
13.2393 | -105.5178
13.243 | -105.522
13.2467 | -105.5246
13.2504 | -105.5225
13.2542 | -105.5181
13.2579 | -105.5105
13.2616 | -105.5027
13.2653 | -105.4923
13.269 | -105.4776
13.2727 | -105.4612
13.2764 | -105.4418
13.2801 | -105.4184
13.2838 | -105.3919
13.2876 | -105.3649
13.2913 | -105.3352
13.295 | -105.3049
13.2987 | -105.2725
13.3024 | -105.2378
13.3061 | -105.2021
13.3098 | -105.1678
13.3135 | -105.1329
13.3172 | -105.0954
13.321 | -105.0571
13.3247 | -105.0224
13.3284 | -104.9902
13.3321 | -104.9579
13.3358 | -104.9274
13.3395 | -104.8993
13.3432 | -104.8758
13.3469 | -104.8556
13.3507 | -104.841
13.3544 | -104.8289
13.3581 | -104.8192
13.3618 | -104.8133
13.3655 | -104.8087
13.3692 | -104.8086
13.3729 | -104.8135
13.3766 | -104.8256
13.3803 | -104.8412
13.3841 | -104.8734
13.3878 | -104.8938
13.3915 | -104.9181
13.3952 | -104.9449
13.3989 | -104.9754
13.4026 | -105.0055
13.4063 | -105.0344
13.41 | -105.0653
13.4138 | -105.0947
13.4175 | -105.1264
13.4212 | -105.1601
13.4249 | -105.1937
13.4286 | -105.2245
13.4323 | -105.255
13.436 | -105.2841
13.4397 | -105.3095
13.4434 | -105.3349
13.4472 | -105.3587
13.4509 | -105.3809
13.4546 | -105.4001
13.4583 | -105.4165
13.462 | -105.428
13.4657 | -105.4387
13.4694 | -105.4491
13.4731 | -105.456
13.4768 | -105.4611
13.4806 | -105.4637
13.4843 | -105.4656
13.488 | -105.4634
13.4917 | -105.4596
13.4954 | -105.4551
13.4991 | -105.4498
13.5028 | -105.4419
13.5065 | -105.4301
13.5103 | -105.4233
13.514 | -105.4217
13.5177 | -105.4196
13.5214 | -105.4166
13.5251 | -105.4164
13.5288 | -105.4166
13.5325 | -105.4163
13.5362 | -105.4149
13.5399 | -105.4173
13.5437 | -105.4197
13.5474 | -105.4249
13.5511 | -105.4328
13.5548 | -105.4374
13.5585 | -105.445
13.5622 | -105.4509
13.5659 | -105.4573
13.5696 | -105.4669
13.5734 | -105.4729
13.5771 | -105.4789
13.5808 | -105.4865
13.5845 | -105.4948
13.5882 | -105.5015
13.5919 | -105.5088
13.5956 | -105.5167
13.5993 | -105.5257
13.603 | -105.5332
13.6068 | -105.5425
13.6105 | -105.5523
13.6142 | -105.56
13.6179 | -105.565
13.6216 | -105.5729
13.6253 | -105.5837
13.629 | -105.5929
13.6327 | -105.5972
13.6364 | -105.6
13.6402 | -105.5991
13.6439 | -105.5989
13.6476 | -105.6016
13.6513 | -105.6022
13.655 | -105.6027
13.6587 | -105.6018
13.6624 | -105.5982
13.6661 | -105.5909
13.6699 | -105.5846
13.6736 | -105.5797
13.6773 | -105.5756
13.681 | -105.5673
13.6847 | -105.5617
13.6884 | -105.5589
13.6921 | -105.5557
13.6958 | -105.5524
13.6995 | -105.5509
13.7033 | -105.5527
13.707 | -105.553
13.7107 | -105.5527
13.7144 | -105.555
13.7181 | -105.5571
13.7218 | -105.5614
13.7255 | -105.5675
13.7292 | -105.5748
13.733 | -105.5809
13.7367 | -105.5864
13.7404 | -105.5922
13.7441 | -105.5949
13.7478 | -105.5997
13.7515 | -105.6116
13.7552 | -105.6201
13.7589 | -105.6251
13.7626 | -105.6266
13.7664 | -105.6233
13.7701 | -105.622
13.7738 | -105.6198
13.7775 | -105.613
13.7812 | -105.6014
13.7849 | -105.5921
13.7886 | -105.579
13.7923 | -105.5629
13.796 | -105.546
13.7998 | -105.5283
13.8035 | -105.5059
13.8072 | -105.4856
13.8109 | -105.4648
13.8146 | -105.4437
13.8183 | -105.4233
13.822 | -105.4047
13.8257 | -105.3869
13.8295 | -105.3761
13.8332 | -105.3704
13.8369 | -105.3666
13.8406 | -105.3624
13.8443 | -105.3574
13.848 | -105.3563
13.8517 | -105.3585
13.8554 | -105.3628
13.8591 | -105.372
13.8629 | -105.3852
13.8666 | -105.3974
13.8703 | -105.41
13.874 | -105.4223
13.8777 | -105.4386
13.8814 | -105.4555
13.8851 | -105.4714
13.8888 | -105.4872
13.8925 | -105.5029
13.8963 | -105.5182
13.9 | -105.5358
13.9037 | -105.5523
13.9074 | -105.5658
13.9111 | -105.5788
13.9148 | -105.5947
13.9185 | -105.6026
13.9222 | -105.616
13.926 | -105.6296
13.9297 | -105.6416
13.9334 | -105.6558
13.9371 | -105.6661
13.9408 | -105.6743
13.9445 | -105.6846
13.9482 | -105.6908
13.9519 | -105.6967
13.9556 | -105.7008
13.9594 | -105.6877
13.9631 | -105.661
13.9668 | -105.6474
13.9705 | -105.6362
13.9742 | -105.6235
13.9779 | -105.6098
13.9816 | -105.5938
13.9853 | -105.5741
13.9891 | -105.5524
13.9928 | -105.5277
13.9965 | -105.4995
14.0002 | -105.4687
14.0039 | -105.4351
14.0076 | -105.3982
14.0113 | -105.3586
14.015 | -105.3176
14.0187 | -105.2752
14.0225 | -105.2345
14.0262 | -105.1966
14.0299 | -105.1567
14.0336 | -105.1168
14.0373 | -105.0818
14.041 | -105.0523
14.0447 | -105.0277
14.0484 | -105.0073
14.0521 | -104.9908
14.0559 | -104.9804
14.0596 | -104.9731
14.0633 | -104.9721
14.067 | -104.9783
14.0707 | -104.9921
14.0744 | -105.0135
14.0781 | -105.0397
14.0818 | -105.0695
14.0856 | -105.1035
14.0893 | -105.1416
14.093 | -105.1819
14.0967 | -105.2228
14.1004 | -105.2646
14.1041 | -105.3062
14.1078 | -105.3526
14.1115 | -105.3975
14.1152 | -105.4387
14.119 | -105.479
14.1227 | -105.5156
14.1264 | -105.5465
14.1301 | -105.575
14.1338 | -105.6015
14.1375 | -105.6227
14.1412 | -105.6415
14.1449 | -105.6587
14.1486 | -105.6691
14.1524 | -105.6763
14.1561 | -105.6747
14.1598 | -105.6708
14.1635 | -105.6797
14.1672 | -105.6911
14.1709 | -105.6825
14.1746 | -105.665
14.1783 | -105.6459
14.1821 | -105.6242
14.1858 | -105.6039
14.1895 | -105.5814
14.1932 | -105.5616
14.1969 | -105.5427
14.2006 | -105.5265
14.2043 | -105.5121
14.208 | -105.4978
14.2117 | -105.4843
14.2155 | -105.4745
14.2192 | -105.4643
14.2229 | -105.4556
14.2266 | -105.4458
14.2303 | -105.4357
14.234 | -105.4306
14.2377 | -105.4297
14.2414 | -105.4284
14.2452 | -105.4287
14.2489 | -105.4315
14.2526 | -105.4344
14.2563 | -105.4419
14.26 | -105.4493
14.2637 | -105.4602
14.2674 | -105.4691
14.2711 | -105.4776
14.2748 | -105.4862
14.2786 | -105.4963
14.2823 | -105.5067
14.286 | -105.5163
14.2897 | -105.5288
14.2934 | -105.5435
14.2971 | -105.5564
14.3008 | -105.5693
14.3045 | -105.5801
14.3082 | -105.5902
14.312 | -105.5957
14.3157 | -105.5999
14.3194 | -105.6053
14.3231 | -105.61
14.3268 | -105.6167
14.3305 | -105.6227
14.3342 | -105.626
14.3379 | -105.6293
14.3417 | -105.6333
14.3454 | -105.6351
14.3491 | -105.6365
14.3528 | -105.6404
14.3565 | -105.6424
14.3602 | -105.6439
14.3639 | -105.6486
14.3676 | -105.6528
14.3713 | -105.6571
14.3751 | -105.6582
14.3788 | -105.6597
14.3825 | -105.6608
14.3862 | -105.6659
14.3899 | -105.6703
14.3936 | -105.6754
14.3973 | -105.6789
14.401 | -105.6801
14.4048 | -105.6795
14.4085 | -105.6796
14.4122 | -105.6792
14.4159 | -105.677
14.4196 | -105.6759
14.4233 | -105.6742
14.427 | -105.6693
14.4307 | -105.6677
14.4344 | -105.6666
14.4382 | -105.6626
14.4419 | -105.6563
14.4456 | -105.6523
14.4493 | -105.6485
14.453 | -105.644
14.4567 | -105.6433
14.4604 | -105.6424
14.4641 | -105.6417
14.4678 | -105.644
14.4716 | -105.644
14.4753 | -105.6466
14.479 | -105.6508
14.4827 | -105.6538
14.4864 | -105.6557
14.4901 | -105.66
14.4938 | -105.6663
14.4975 | -105.67
14.5013 | -105.6712
14.505 | -105.6742
14.5087 | -105.6773
14.5124 | -105.6807
14.5161 | -105.6862
14.5198 | -105.6901
14.5235 | -105.6927
14.5272 | -105.6957
14.5309 | -105.6983
14.5347 | -105.6958
14.5384 | -105.6965
14.5421 | -105.6974
14.5458 | -105.6952
14.5495 | -105.6918
14.5532 | -105.6886
14.5569 | -105.6833
14.5606 | -105.6777
14.5643 | -105.6742
14.5681 | -105.6751
14.5718 | -105.6726
14.5755 | -105.6668
14.5792 | -105.6614
14.5829 | -105.6575
14.5866 | -105.6527
14.5903 | -105.6497
14.594 | -105.6447
14.5978 | -105.643
14.6015 | -105.6413
14.6052 | -105.6427
14.6089 | -105.6482
14.6126 | -105.6539
14.6163 | -105.6592
14.62 | -105.6685
14.6237 | -105.678
14.6274 | -105.6876
14.6312 | -105.6998
14.6349 | -105.7152
14.6386 | -105.7294
14.6423 | -105.7432
14.646 | -105.7583
14.6497 | -105.7755
14.6534 | -105.7912
14.6571 | -105.8084
14.6609 | -105.8257
14.6646 | -105.8405
14.6683 | -105.8559
14.672 | -105.8679
14.6757 | -105.88
14.6794 | -105.8921
14.6831 | -105.9054
14.6868 | -105.9179
14.6905 | -105.9279
14.6943 | -105.9365
14.698 | -105.945
14.7017 | -105.9494
14.7054 | -105.9537
14.7091 | -105.9549
14.7128 | -105.9564
14.7165 | -105.9584
14.7202 | -105.9582
14.7239 | -105.9504
14.7277 | -105.937
14.7314 | -105.9267
14.7351 | -105.9161
14.7388 | -105.8989
14.7425 | -105.8842
14.7462 | -105.8735
14.7499 | -105.859
14.7536 | -105.8462
14.7574 | -105.8351
14.7611 | -105.8243
14.7648 | -105.8142
14.7685 | -105.8053
14.7722 | -105.7975
14.7759 | -105.7893
14.7796 | -105.7823
14.7833 | -105.7772
14.787 | -105.7761
14.7908 | -105.7766
14.7945 | -105.7794
14.7982 | -105.7852
14.8019 | -105.7893
14.8056 | -105.7946
14.8093 | -105.8008
14.813 | -105.8098
14.8167 | -105.8185
14.8205 | -105.8263
14.8242 | -105.8368
14.8279 | -105.8481
14.8316 | -105.8547
14.8353 | -105.8638
14.839 | -105.8715
14.8427 | -105.8791
14.8464 | -105.8872
14.8501 | -105.8946
14.8539 | -105.9012
14.8576 | -105.9052
14.8613 | -105.9076
14.865 | -105.9079
14.8687 | -105.9049
14.8724 | -105.9038
14.8761 | -105.9033
14.8798 | -105.9009
14.8835 | -105.8969
14.8873 | -105.8893
14.891 | -105.8804
14.8947 | -105.8687
14.8984 | -105.8563
14.9021 | -105.8452
14.9058 | -105.832
14.9095 | -105.8174
14.9132 | -105.8033
14.917 | -105.785
14.9207 | -105.7679
14.9244 | -105.7496
14.9281 | -105.7291
14.9318 | -105.7092
14.9355 | -105.6911
14.9392 | -105.6749
14.9429 | -105.6569
14.9466 | -105.6372
14.9504 | -105.619
14.9541 | -105.5986
14.9578 | -105.5791
14.9615 | -105.56
14.9652 | -105.5386
14.9689 | -105.5208
14.9726 | -105.5035
14.9763 | -105.4862
14.9801 | -105.473
14.9838 | -105.4574
14.9875 | -105.4368
14.9912 | -105.4159
14.9949 | -105.4026
14.9986 | -105.391
15.0023 | -105.3784
15.006 | -105.3659
15.0097 | -105.3543
15.0135 | -105.3447
15.0172 | -105.3373
15.0209 | -105.3324
15.0246 | -105.3239
15.0283 | -105.3155
15.032 | -105.3073
15.0357 | -105.301
15.0394 | -105.2956
15.0431 | -105.2906
15.0469 | -105.2892
15.0506 | -105.2877
15.0543 | -105.2874
15.058 | -105.287
15.0617 | -105.2839
15.0654 | -105.2806
15.0691 | -105.2819
15.0728 | -105.2841
15.0766 | -105.2852
15.0803 | -105.2891
15.084 | -105.2936
15.0877 | -105.2957
15.0914 | -105.2993
15.0951 | -105.3046
15.0988 | -105.3123
15.1025 | -105.3202
15.1062 | -105.3282
15.11 | -105.3365
15.1137 | -105.3476
15.1174 | -105.3599
15.1211 | -105.3729
15.1248 | -105.387
15.1285 | -105.4031
15.1322 | -105.4212
15.1359 | -105.4401
15.1396 | -105.4591
15.1434 | -105.4766
15.1471 | -105.4954
15.1508 | -105.5181
15.1545 | -105.5402
15.1582 | -105.5604
15.1619 | -105.5815
15.1656 | -105.6023
15.1693 | -105.6211
15.1731 | -105.6389
15.1768 | -105.6565
15.1805 | -105.6772
15.1842 | -105.6945
15.1879 | -105.7108
15.1916 | -105.7271
15.1953 | -105.7422
15.199 | -105.7599
15.2027 | -105.7801
15.2065 | -105.7964
15.2102 | -105.8075
15.2139 | -105.8153
15.2176 | -105.8217
15.2213 | -105.8279
15.225 | -105.8336
15.2287 | -105.8376
15.2324 | -105.8398
15.2362 | -105.8405
15.2399 | -105.84
15.2436 | -105.8409
15.2473 | -105.8412
15.251 | -105.8421
15.2547 | -105.8418
15.2584 | -105.8395
15.2621 | -105.8379
15.2658 | -105.8376
15.2696 | -105.839
15.2733 | -105.835
15.277 | -105.8305
15.2807 | -105.8301
15.2844 | -105.8304
15.2881 | -105.8277
15.2918 | -105.827
15.2955 | -105.8269
15.2992 | -105.8238
15.303 | -105.8185
15.3067 | -105.8129
15.3104 | -105.804
15.3141 | -105.795
15.3178 | -105.7856
15.3215 | -105.7756
15.3252 | -105.7687
15.3289 | -105.759
15.3327 | -105.7449
15.3364 | -105.7305
15.3401 | -105.7169
15.3438 | -105.7039
15.3475 | -105.6891
15.3512 | -105.6746
15.3549 | -105.6598
15.3586 | -105.647
15.3623 | -105.6357
15.3661 | -105.622
15.3698 | -105.607
15.3735 | -105.5932
15.3772 | -105.5826
15.3809 | -105.5716
15.3846 | -105.5631
15.3883 | -105.5561
15.392 | -105.5479
15.3957 | -105.5423
15.3995 | -105.5387
15.4032 | -105.534
15.4069 | -105.5309
15.4106 | -105.529
15.4143 | -105.5259
15.418 | -105.5247
15.4217 | -105.5229
15.4254 | -105.5255
15.4292 | -105.5291
15.4329 | -105.5327
15.4366 | -105.5369
15.4403 | -105.5413
15.444 | -105.5464
15.4477 | -105.5493
15.4514 | -105.553
15.4551 | -105.5547
15.4588 | -105.5564
15.4626 | -105.5602
15.4663 | -105.5625
15.47 | -105.5635
15.4737 | -105.5643
15.4774 | -105.5649
15.4811 | -105.5638
15.4848 | -105.5602
15.4885 | -105.5547
15.4923 | -105.552
15.496 | -105.5494
15.4997 | -105.5445
15.5034 | -105.5393
15.5071 | -105.533
15.5108 | -105.5259
15.5145 | -105.5177
15.5182 | -105.5096
15.5219 | -105.5054
15.5257 | -105.4998
15.5294 | -105.4942
15.5331 | -105.4888
15.5368 | -105.482
15.5405 | -105.4766
15.5442 | -105.4704
15.5479 | -105.4626
15.5516 | -105.4598
15.5553 | -105.4566
15.5591 | -105.4547
15.5628 | -105.4539
15.5665 | -105.4513
15.5702 | -105.4473
15.5739 | -105.4421
15.5776 | -105.4393
15.5813 | -105.4365
15.585 | -105.4322
15.5888 | -105.4293
15.5925 | -105.4285
15.5962 | -105.4273
15.5999 | -105.4218
15.6036 | -105.4195
15.6073 | -105.4172
15.611 | -105.4122
15.6147 | -105.4074
15.6184 | -105.3991
15.6222 | -105.3883
15.6259 | -105.3819
15.6296 | -105.3757
15.6333 | -105.3683
15.637 | -105.3632
15.6407 | -105.3584
15.6444 | -105.3557
15.6481 | -105.3518
15.6519 | -105.3463
15.6556 | -105.3406
15.6593 | -105.3363
15.663 | -105.3325
15.6667 | -105.3286
15.6704 | -105.3276
15.6741 | -105.3261
15.6778 | -105.3285
15.6815 | -105.3299
15.6853 | -105.3302
15.689 | -105.3331
15.6927 | -105.3372
15.6964 | -105.3407
15.7001 | -105.3452
15.7038 | -105.3524
15.7075 | -105.3613
15.7112 | -105.37
15.7149 | -105.381
15.7187 | -105.3931
15.7224 | -105.4064
15.7261 | -105.4193
15.7298 | -105.4327
15.7335 | -105.4459
15.7372 | -105.4608
15.7409 | -105.476
15.7446 | -105.489
15.7484 | -105.5046
15.7521 | -105.5209
15.7558 | -105.5356
15.7595 | -105.5497
15.7632 | -105.5626
15.7669 | -105.5731
15.7706 | -105.5846
15.7743 | -105.5952
15.778 | -105.6023
15.7818 | -105.608
15.7855 | -105.6112
15.7892 | -105.6133
15.7929 | -105.6163
15.7966 | -105.6204
15.8003 | -105.6232
15.804 | -105.6218
15.8077 | -105.6168
15.8115 | -105.6132
15.8152 | -105.609
15.8189 | -105.6031
15.8226 | -105.5955
15.8263 | -105.5868
15.83 | -105.5759
15.8337 | -105.5642
15.8374 | -105.5551
15.8411 | -105.5458
15.8449 | -105.5367
15.8486 | -105.5252
15.8523 | -105.5105
15.856 | -105.4956
15.8597 | -105.4813
15.8634 | -105.464
15.8671 | -105.4465
15.8708 | -105.4317
15.8745 | -105.4174
15.8783 | -105.402
15.882 | -105.3873
15.8857 | -105.376
15.8894 | -105.3631
15.8931 | -105.3494
15.8968 | -105.3359
15.9005 | -105.3222
15.9042 | -105.3085
15.908 | -105.2948
15.9117 | -105.2831
15.9154 | -105.2728
15.9191 | -105.2618
15.9228 | -105.2554
15.9265 | -105.2512
15.9302 | -105.2452
15.9339 | -105.2384
15.9376 | -105.2329
15.9414 | -105.227
15.9451 | -105.2213
15.9488 | -105.2144
15.9525 | -105.2105
15.9562 | -105.2081
15.9599 | -105.2076
15.9636 | -105.2064
15.9673 | -105.2058
15.971 | -105.2057
15.9748 | -105.2033
15.9785 | -105.2012
15.9822 | -105.2014
15.9859 | -105.203
15.9896 | -105.2048
15.9933 | -105.2067
15.997 | -105.2087
16.0007 | -105.212
16.0045 | -105.2178
16.0082 | -105.2234
16.0119 | -105.2294
16.0156 | -105.2361
16.0193 | -105.2426
16.023 | -105.2486
16.0267 | -105.2545
16.0304 | -105.2635
16.0341 | -105.2741
16.0379 | -105.2835
16.0416 | -105.293
16.0453 | -105.3018
16.049 | -105.3137
16.0527 | -105.3289
16.0564 | -105.3427
16.0601 | -105.3568
16.0638 | -105.3769
16.0676 | -105.3941
16.0713 | -105.4122
16.075 | -105.4297
16.0787 | -105.4473
16.0824 | -105.4677
16.0861 | -105.4886
16.0898 | -105.5067
16.0935 | -105.5278
16.0972 | -105.5481
16.101 | -105.566
16.1047 | -105.5845
16.1084 | -105.6049
16.1121 | -105.6224
16.1158 | -105.6376
16.1195 | -105.6569
16.1232 | -105.6734
16.1269 | -105.6882
16.1306 | -105.7029
16.1344 | -105.7137
16.1381 | -105.7236
16.1418 | -105.7328
16.1455 | -105.7396
16.1492 | -105.7436
16.1529 | -105.7419
16.1566 | -105.7405
16.1603 | -105.7417
16.1641 | -105.7421
16.1678 | -105.7414
16.1715 | -105.7407
16.1752 | -105.7375
16.1789 | -105.7343
16.1826 | -105.7265
16.1863 | -105.7179
16.19 | -105.7093
16.1937 | -105.7015
16.1975 | -105.6936
16.2012 | -105.6845
16.2049 | -105.6756
16.2086 | -105.6692
16.2123 | -105.6625
16.216 | -105.6517
16.2197 | -105.641
16.2234 | -105.6307
16.2272 | -105.623
16.2309 | -105.6158
16.2346 | -105.6073
16.2383 | -105.6031
16.242 | -105.5997
16.2457 | -105.595
16.2494 | -105.5927
16.2531 | -105.5919
16.2568 | -105.5921
16.2606 | -105.5911
16.2643 | -105.5913
16.268 | -105.5907
16.2717 | -105.5923
16.2754 | -105.5955
16.2791 | -105.5968
16.2828 | -105.6005
16.2865 | -105.6066
16.2902 | -105.6117
16.294 | -105.6194
16.2977 | -105.6238
16.3014 | -105.6274
16.3051 | -105.6305
16.3088 | -105.6351
16.3125 | -105.6423
16.3162 | -105.6481
16.3199 | -105.6521
16.3237 | -105.6579
16.3274 | -105.6581
16.3311 | -105.6591
16.3348 | -105.6621
16.3385 | -105.6627
16.3422 | -105.6618
16.3459 | -105.6628
16.3496 | -105.6622
16.3533 | -105.6592
16.3571 | -105.657
16.3608 | -105.6542
16.3645 | -105.6506
16.3682 | -105.6479
16.3719 | -105.6476
16.3756 | -105.6464
16.3793 | -105.646
16.383 | -105.6497
16.3867 | -105.6555
16.3905 | -105.657
16.3942 | -105.6578
16.3979 | -105.6621
16.4016 | -105.67
16.4053 | -105.6779
16.409 | -105.6874
16.4127 | -105.6969
16.4164 | -105.711
16.4202 | -105.7257
16.4239 | -105.7396
16.4276 | -105.7535
16.4313 | -105.7676
16.435 | -105.7816
16.4387 | -105.7981
16.4424 | -105.8142
16.4461 | -105.8281
16.4498 | -105.8447
16.4536 | -105.8619
16.4573 | -105.8785
16.461 | -105.8914
16.4647 | -105.9011
16.4684 | -105.9112
16.4721 | -105.9183
16.4758 | -105.9222
16.4795 | -105.9276
16.4833 | -105.9318
16.487 | -105.9361
16.4907 | -105.9401
16.4944 | -105.9419
16.4981 | -105.9436
16.5018 | -105.9471
16.5055 | -105.9447
16.5092 | -105.9431
16.5129 | -105.9427
16.5167 | -105.9416
16.5204 | -105.9407
16.5241 | -105.9398
16.5278 | -105.9378
16.5315 | -105.9395
16.5352 | -105.9425
16.5389 | -105.9473
16.5426 | -105.9508
16.5463 | -105.953
16.5501 | -105.9559
16.5538 | -105.9597
16.5575 | -105.9636
16.5612 | -105.9718
16.5649 | -105.9785
16.5686 | -105.9842
16.5723 | -105.9928
16.576 | -105.9984
16.5798 | -106.0018
16.5835 | -106.0034
16.5872 | -106.0081
16.5909 | -106.0148
16.5946 | -106.0189
16.5983 | -106.0174
16.602 | -106.0187
16.6057 | -106.0201
16.6094 | -106.0189
16.6132 | -106.0162
16.6169 | -106.0138
16.6206 | -106.0109
16.6243 | -106.0054
16.628 | -106.0021
16.6317 | -105.9976
16.6354 | -105.9904
16.6391 | -105.9825
16.6429 | -105.975
16.6466 | -105.9659
16.6503 | -105.96
16.654 | -105.9525
16.6577 | -105.9453
16.6614 | -105.939
16.6651 | -105.9338
16.6688 | -105.9297
16.6725 | -105.9265
16.6763 | -105.9208
16.68 | -105.915
16.6837 | -105.9129
16.6874 | -105.9109
16.6911 | -105.9092
16.6948 | -105.9076
16.6985 | -105.9099
16.7022 | -105.9156
16.7059 | -105.9201
16.7097 | -105.9257
16.7134 | -105.9341
16.7171 | -105.9393
16.7208 | -105.9432
16.7245 | -105.9467
16.7282 | -105.9528
16.7319 | -105.9592
16.7356 | -105.9649
16.7394 | -105.9744
16.7431 | -105.9868
16.7468 | -105.9981
16.7505 | -106.0116
16.7542 | -106.0227
16.7579 | -106.0332
16.7616 | -106.0465
16.7653 | -106.0599
16.769 | -106.0705
16.7728 | -106.0821
16.7765 | -106.0935
16.7802 | -106.1067
16.7839 | -106.12
16.7876 | -106.1345
16.7913 | -106.1507
16.795 | -106.1682
16.7987 | -106.183
16.8025 | -106.1973
16.8062 | -106.2087
16.8099 | -106.2174
16.8136 | -106.2272
16.8173 | -106.2382
16.821 | -106.2484
16.8247 | -106.2536
16.8284 | -106.2571
16.8321 | -106.2611
16.8359 | -106.2659
16.8396 | -106.2668
16.8433 | -106.267
16.847 | -106.2667
16.8507 | -106.2671
16.8544 | -106.2677
16.8581 | -106.2666
16.8618 | -106.2624
16.8655 | -106.2593
16.8693 | -106.2534
16.873 | -106.2441
16.8767 | -106.234
16.8804 | -106.2237
16.8841 | -106.2124
16.8878 | -106.203
16.8915 | -106.189
16.8952 | -106.1728
16.899 | -106.157
16.9027 | -106.1404
16.9064 | -106.1232
16.9101 | -106.1072
16.9138 | -106.0925
16.9175 | -106.0777
16.9212 | -106.0631
16.9249 | -106.0463
16.9286 | -106.03
16.9324 | -106.0151
16.9361 | -106.0006
16.9398 | -105.9837
16.9435 | -105.9676
16.9472 | -105.9549
16.9509 | -105.9425
16.9546 | -105.929
16.9583 | -105.9157
16.962 | -105.9051
16.9658 | -105.8967
16.9695 | -105.8868
16.9732 | -105.8755
16.9769 | -105.8661
16.9806 | -105.8542
16.9843 | -105.8433
16.988 | -105.8346
16.9917 | -105.826
16.9955 | -105.8204
16.9992 | -105.8151
17.0029 | -105.8089
17.0066 | -105.805
17.0103 | -105.8046
17.014 | -105.8031
17.0177 | -105.8028
17.0214 | -105.8027
17.0251 | -105.8022
17.0289 | -105.8018
17.0326 | -105.8012
17.0363 | -105.8003
17.04 | -105.8029
17.0437 | -105.8082
17.0474 | -105.8163
17.0511 | -105.8246
17.0548 | -105.8321
17.0586 | -105.8397
17.0623 | -105.8479
17.066 | -105.8553
17.0697 | -105.863
17.0734 | -105.8731
17.0771 | -105.8805
17.0808 | -105.8901
17.0845 | -105.9014
17.0882 | -105.9111
17.092 | -105.9192
17.0957 | -105.9266
17.0994 | -105.9349
17.1031 | -105.9442
17.1068 | -105.9543
17.1105 | -105.9646
17.1142 | -105.9774
17.1179 | -105.9883
17.1216 | -105.9961
17.1254 | -106.0017
17.1291 | -106.0067
17.1328 | -106.0065
17.1365 | -106.0108
17.1402 | -106.0154
17.1439 | -106.0187
17.1476 | -106.0216
17.1513 | -106.0236
17.1551 | -106.0213
17.1588 | -106.0181
17.1625 | -106.0153
17.1662 | -106.0125
17.1699 | -106.0093
17.1736 | -106.0048
17.1773 | -106.0001
17.181 | -105.9944
17.1847 | -105.9893
17.1885 | -105.9832
17.1922 | -105.978
17.1959 | -105.9732
17.1996 | -105.971
17.2033 | -105.9696
17.207 | -105.9637
17.2107 | -105.9566
17.2144 | -105.9504
17.2182 | -105.9482
17.2219 | -105.9459
17.2256 | -105.9438
17.2293 | -105.9419
17.233 | -105.9399
17.2367 | -105.9395
17.2404 | -105.9395
17.2441 | -105.9375
17.2478 | -105.938
17.2516 | -105.9385
17.2553 | -105.9368
17.259 | -105.9367
17.2627 | -105.9361
17.2664 | -105.9349
17.2701 | -105.9353
17.2738 | -105.9355
17.2775 | -105.9377
17.2812 | -105.9403
17.285 | -105.9412
17.2887 | -105.9416
17.2924 | -105.9388
17.2961 | -105.9362
17.2998 | -105.9374
17.3035 | -105.937
17.3072 | -105.9346
17.3109 | -105.9318
17.3147 | -105.9287
17.3184 | -105.9226
17.3221 | -105.9175
17.3258 | -105.9131
17.3295 | -105.9079
17.3332 | -105.9029
17.3369 | -105.9006
17.3406 | -105.8958
17.3443 | -105.8895
17.3481 | -105.8858
17.3518 | -105.8807
17.3555 | -105.879
17.3592 | -105.8754
17.3629 | -105.8735
17.3666 | -105.8701
17.3703 | -105.8679
17.374 | -105.8632
17.3777 | -105.8608
17.3815 | -105.8585
17.3852 | -105.8564
17.3889 | -105.8542
17.3926 | -105.8526
17.3963 | -105.8513
17.4 | -105.8505
17.4037 | -105.8501
17.4074 | -105.8501
17.4112 | -105.8482
17.4149 | -105.8461
17.4186 | -105.8447
17.4223 | -105.841
17.426 | -105.8373
17.4297 | -105.8345
17.4334 | -105.8307
17.4371 | -105.8271
17.4408 | -105.8236
17.4446 | -105.8183
17.4483 | -105.8128
17.452 | -105.8093
17.4557 | -105.8075
17.4594 | -105.8072
17.4631 | -105.8058
17.4668 | -105.8028
17.4705 | -105.7999
17.4743 | -105.7966
17.478 | -105.7933
17.4817 | -105.7908
17.4854 | -105.7886
17.4891 | -105.7846
17.4928 | -105.7834
17.4965 | -105.7814
17.5002 | -105.7782
17.5039 | -105.7761
17.5077 | -105.7711
17.5114 | -105.7667
17.5151 | -105.7634
17.5188 | -105.7576
17.5225 | -105.7508
17.5262 | -105.7413
17.5299 | -105.7314
17.5336 | -105.7252
17.5373 | -105.7179
17.5411 | -105.7135
17.5448 | -105.7086
17.5485 | -105.7017
17.5522 | -105.6959
17.5559 | -105.6928
17.5596 | -105.6876
17.5633 | -105.6838
17.567 | -105.6819
17.5708 | -105.6797
17.5745 | -105.6781
17.5782 | -105.6774
17.5819 | -105.6773
17.5856 | -105.6763
17.5893 | -105.6786
17.593 | -105.6802
17.5967 | -105.6799
17.6004 | -105.6801
17.6042 | -105.6795
17.6079 | -105.6792
17.6116 | -105.6805
17.6153 | -105.6854
17.619 | -105.6888
17.6227 | -105.6934
17.6264 | -105.6963
17.6301 | -105.6972
17.6338 | -105.6988
17.6376 | -105.7018
17.6413 | -105.7037
17.645 | -105.7038
17.6487 | -105.7022
17.6524 | -105.7036
17.6561 | -105.7061
17.6598 | -105.71
17.6635 | -105.7105
17.6673 | -105.7077
17.671 | -105.7073
17.6747 | -105.7109
17.6784 | -105.7122
17.6821 | -105.712
17.6858 | -105.7129
17.6895 | -105.7146
17.6932 | -105.7147
17.6969 | -105.7162
17.7007 | -105.7168
17.7044 | -105.7194
17.7081 | -105.7214
17.7118 | -105.7231
17.7155 | -105.7246
17.7192 | -105.7265
17.7229 | -105.7274
17.7266 | -105.7288
17.7304 | -105.731
17.7341 | -105.7297
17.7378 | -105.7257
17.7415 | -105.7266
17.7452 | -105.7282
17.7489 | -105.7272
17.7526 | -105.7256
17.7563 | -105.7252
17.76 | -105.7274
17.7638 | -105.727
17.7675 | -105.7248
17.7712 | -105.7235
17.7749 | -105.721
17.7786 | -105.7178
17.7823 | -105.7165
17.786 | -105.7132
17.7897 | -105.7085
17.7934 | -105.7033
17.7972 | -105.6987
17.8009 | -105.6943
17.8046 | -105.6904
17.8083 | -105.6858
17.812 | -105.683
17.8157 | -105.6768
17.8194 | -105.672
17.8231 | -105.6692
17.8269 | -105.6654
17.8306 | -105.6623
17.8343 | -105.66
17.838 | -105.6561
17.8417 | -105.6533
17.8454 | -105.652
17.8491 | -105.65
17.8528 | -105.6489
17.8565 | -105.6494
17.8603 | -105.6507
17.864 | -105.6509
17.8677 | -105.6498
17.8714 | -105.6503
17.8751 | -105.6518
17.8788 | -105.6522
17.8825 | -105.6536
17.8862 | -105.6542
17.89 | -105.6546
17.8937 | -105.6578
17.8974 | -105.6624
17.9011 | -105.6646
17.9048 | -105.6647
17.9085 | -105.6638
17.9122 | -105.6622
17.9159 | -105.6626
17.9196 | -105.6666
17.9234 | -105.6685
17.9271 | -105.6707
17.9308 | -105.6705
17.9345 | -105.6711
17.9382 | -105.6722
17.9419 | -105.671
17.9456 | -105.6688
17.9493 | -105.6667
17.953 | -105.6647
17.9568 | -105.6632
17.9605 | -105.6584
17.9642 | -105.6548
17.9679 | -105.6495
17.9716 | -105.6463
17.9753 | -105.6428
17.979 | -105.6388
17.9827 | -105.6346
17.9865 | -105.6278
17.9902 | -105.6248
17.9939 | -105.6244
17.9976 | -105.6199
18.0013 | -105.612
18.005 | -105.6039
18.0087 | -105.5978
18.0124 | -105.5923
18.0161 | -105.585
18.0199 | -105.5776
18.0236 | -105.5718
18.0273 | -105.5658
18.031 | -105.5608
18.0347 | -105.5573
18.0384 | -105.5497
18.0421 | -105.543
18.0458 | -105.5363
18.0495 | -105.5299
18.0533 | -105.5219
18.057 | -105.5146
18.0607 | -105.5062
18.0644 | -105.4986
18.0681 | -105.4929
18.0718 | -105.4871
18.0755 | -105.4801
18.0792 | -105.4713
18.083 | -105.463
18.0867 | -105.4546
18.0904 | -105.4445
18.0941 | -105.4375
18.0978 | -105.4296
18.1015 | -105.4215
18.1052 | -105.4144
18.1089 | -105.4078
18.1126 | -105.3985
18.1164 | -105.3901
18.1201 | -105.3818
18.1238 | -105.3767
18.1275 | -105.3727
18.1312 | -105.3683
18.1349 | -105.3637
18.1386 | -105.3596
18.1423 | -105.3557
18.1461 | -105.3522
18.1498 | -105.3483
18.1535 | -105.3453
18.1572 | -105.3462
18.1609 | -105.3472
18.1646 | -105.3459
18.1683 | -105.3446
18.172 | -105.3441
18.1757 | -105.3424
18.1795 | -105.3418
18.1832 | -105.3405
18.1869 | -105.3399
18.1906 | -105.3389
18.1943 | -105.3364
18.198 | -105.3322
18.2017 | -105.3302
18.2054 | -105.325
18.2091 | -105.324
18.2129 | -105.3227
18.2166 | -105.3184
18.2203 | -105.3149
18.224 | -105.3106
18.2277 | -105.3063
18.2314 | -105.299
18.2351 | -105.2922
18.2388 | -105.2881
18.2426 | -105.2842
18.2463 | -105.2801
18.25 | -105.2737
18.2537 | -105.2651
18.2574 | -105.259
18.2611 | -105.255
18.2648 | -105.252
18.2685 | -105.248
18.2722 | -105.2423
18.276 | -105.2363
18.2797 | -105.2307
18.2834 | -105.2247
18.2871 | -105.2197
18.2908 | -105.2155
18.2945 | -105.2097
18.2982 | -105.2048
18.3019 | -105.1996
18.3057 | -105.1962
18.3094 | -105.194
18.3131 | -105.1925
18.3168 | -105.1899
18.3205 | -105.1881
18.3242 | -105.1859
18.3279 | -105.1857
18.3316 | -105.1859
18.3353 | -105.1849
18.3391 | -105.1888
18.3428 | -105.1948
18.3465 | -105.2001
18.3502 | -105.2058
18.3539 | -105.2135
18.3576 | -105.2221
18.3613 | -105.2292
18.365 | -105.2375
18.3687 | -105.2462
18.3725 | -105.2543
18.3762 | -105.2617
18.3799 | -105.2701
18.3836 | -105.2773
18.3873 | -105.2843
18.391 | -105.2917
18.3947 | -105.2993
18.3984 | -105.3054
18.4022 | -105.3126
18.4059 | -105.3183
18.4096 | -105.3245
18.4133 | -105.3307
18.417 | -105.3332
18.4207 | -105.3349
18.4244 | -105.3402
18.4281 | -105.3467
18.4318 | -105.3498
18.4356 | -105.3493
18.4393 | -105.3488
18.443 | -105.3484
18.4467 | -105.3481
18.4504 | -105.3478
18.4541 | -105.35
18.4578 | -105.3527
18.4615 | -105.3503
18.4653 | -105.3489
18.469 | -105.3463
18.4727 | -105.34
18.4764 | -105.333
18.4801 | -105.3276
18.4838 | -105.3216
18.4875 | -105.3161
18.4912 | -105.3098
18.4949 | -105.3062
18.4987 | -105.3001
18.5024 | -105.2908
18.5061 | -105.2831
18.5098 | -105.2768
18.5135 | -105.2704
18.5172 | -105.262
18.5209 | -105.2506
18.5246 | -105.2404
18.5283 | -105.2305
18.5321 | -105.2219
18.5358 | -105.213
18.5395 | -105.2048
18.5432 | -105.1978
18.5469 | -105.1911
18.5506 | -105.1829
18.5543 | -105.1747
18.558 | -105.17
18.5618 | -105.1666
18.5655 | -105.1628
18.5692 | -105.16
18.5729 | -105.1559
18.5766 | -105.1535
18.5803 | -105.1515
18.584 | -105.1489
18.5877 | -105.1453
18.5914 | -105.1387
18.5952 | -105.1347
18.5989 | -105.1343
18.6026 | -105.1319
18.6063 | -105.1298
18.61 | -105.1296
18.6137 | -105.1286
18.6174 | -105.1242
18.6211 | -105.1229
18.6248 | -105.1214
18.6286 | -105.1196
18.6323 | -105.1194
18.636 | -105.1206
18.6397 | -105.1223
18.6434 | -105.1217
18.6471 | -105.1225
18.6508 | -105.1243
18.6545 | -105.1266
18.6583 | -105.1282
18.662 | -105.1315
18.6657 | -105.1344
18.6694 | -105.1366
18.6731 | -105.1361
18.6768 | -105.1372
18.6805 | -105.1434
18.6842 | -105.1496
18.6879 | -105.1534
18.6917 | -105.1536
18.6954 | -105.1572
18.6991 | -105.1641
18.7028 | -105.1705
18.7065 | -105.1782
18.7102 | -105.1836
18.7139 | -105.1912
18.7176 | -105.1986
18.7214 | -105.2071
18.7251 | -105.215
18.7288 | -105.2236
18.7325 | -105.2316
18.7362 | -105.2404
18.7399 | -105.2452
18.7436 | -105.2529
18.7473 | -105.2598
18.751 | -105.2653
18.7548 | -105.2703
18.7585 | -105.2755
18.7622 | -105.276
18.7659 | -105.2752
18.7696 | -105.275
18.7733 | -105.2714
18.777 | -105.2691
18.7807 | -105.2677
18.7844 | -105.2662
18.7882 | -105.2645
18.7919 | -105.2601
18.7956 | -105.2549
18.7993 | -105.2473
18.803 | -105.2399
18.8067 | -105.2309
18.8104 | -105.2194
18.8141 | -105.2068
18.8179 | -105.1951
18.8216 | -105.1843
18.8253 | -105.1704
18.829 | -105.1541
18.8327 | -105.1385
18.8364 | -105.1229
18.8401 | -105.1098
18.8438 | -105.096
18.8475 | -105.0829
18.8513 | -105.0689
18.855 | -105.0569
18.8587 | -105.0459
18.8624 | -105.0331
18.8661 | -105.0165
18.8698 | -104.9975
18.8735 | -104.9794
18.8772 | -104.9606
18.8809 | -104.9476
18.8847 | -104.9354
18.8884 | -104.9224
18.8921 | -104.9111
18.8958 | -104.898
18.8995 | -104.8839
18.9032 | -104.8722
18.9069 | -104.8614
18.9106 | -104.8503
18.9144 | -104.8368
18.9181 | -104.8264
18.9218 | -104.8181
18.9255 | -104.8102
18.9292 | -104.8018
18.9329 | -104.7917
18.9366 | -104.7825
18.9403 | -104.776
18.944 | -104.7695
18.9478 | -104.7622
18.9515 | -104.7543
18.9552 | -104.7486
18.9589 | -104.7465
18.9626 | -104.7438
18.9663 | -104.7394
18.97 | -104.7334
18.9737 | -104.7274
18.9775 | -104.7238
18.9812 | -104.7185
18.9849 | -104.7144
18.9886 | -104.7125
18.9923 | -104.7119
18.996 | -104.709
18.9997 | -104.7053
19.0034 | -104.703
19.0071 | -104.7023
19.0109 | -104.7032
19.0146 | -104.7022
19.0183 | -104.7007
19.022 | -104.7027
19.0257 | -104.7041
19.0294 | -104.7054
19.0331 | -104.7089
19.0368 | -104.7145
19.0405 | -104.7196
19.0443 | -104.725
19.048 | -104.7282
19.0517 | -104.73
19.0554 | -104.7332
19.0591 | -104.7373
19.0628 | -104.7387
19.0665 | -104.7412
19.0702 | -104.7448
19.074 | -104.7495
19.0777 | -104.7549
19.0814 | -104.7621
19.0851 | -104.7686
19.0888 | -104.7737
19.0925 | -104.7781
19.0962 | -104.7838
19.0999 | -104.7906
19.1036 | -104.7953
19.1074 | -104.7997
19.1111 | -104.803
19.1148 | -104.8055
19.1185 | -104.8061
19.1222 | -104.8095
19.1259 | -104.8135
19.1296 | -104.816
19.1333 | -104.8179
19.1371 | -104.8213
19.1408 | -104.8257
19.1445 | -104.829
19.1482 | -104.8317
19.1519 | -104.8314
19.1556 | -104.8296
19.1593 | -104.8216
19.163 | -104.816
19.1667 | -104.8132
19.1705 | -104.8101
19.1742 | -104.804
19.1779 | -104.7976
19.1816 | -104.7938
19.1853 | -104.7907
19.189 | -104.781
19.1927 | -104.7687
19.1964 | -104.7596
19.2001 | -104.7536
19.2039 | -104.748
19.2076 | -104.7414
19.2113 | -104.7339
19.215 | -104.7266
19.2187 | -104.7174
19.2224 | -104.7103
19.2261 | -104.7063
19.2298 | -104.7015
19.2336 | -104.6957
19.2373 | -104.6907
19.241 | -104.6888
19.2447 | -104.6879
19.2484 | -104.6872
19.2521 | -104.6876
19.2558 | -104.6874
19.2595 | -104.6854
19.2632 | -104.6836
19.267 | -104.684
19.2707 | -104.6891
19.2744 | -104.6946
19.2781 | -104.6974
19.2818 | -104.6994
19.2855 | -104.7027
19.2892 | -104.7065
19.2929 | -104.7122
19.2966 | -104.7154
19.3004 | -104.7182
19.3041 | -104.722
19.3078 | -104.7258
19.3115 | -104.7312
19.3152 | -104.7377
19.3189 | -104.7448
19.3226 | -104.7526
19.3263 | -104.761
19.3301 | -104.7687
19.3338 | -104.7761
19.3375 | -104.7862
19.3412 | -104.7945
19.3449 | -104.7998
19.3486 | -104.8051
19.3523 | -104.8101
19.356 | -104.8133
19.3597 | -104.8176
19.3635 | -104.8227
19.3672 | -104.8269
19.3709 | -104.83
19.3746 | -104.8334
19.3783 | -104.8366
19.382 | -104.8378
19.3857 | -104.8387
19.3894 | -104.8393
19.3932 | -104.8418
19.3969 | -104.8421
19.4006 | -104.8399
19.4043 | -104.8379
19.408 | -104.8368
19.4117 | -104.8344
19.4154 | -104.8309
19.4191 | -104.8254
19.4228 | -104.8203
19.4266 | -104.813
19.4303 | -104.806
19.434 | -104.8058
19.4377 | -104.8005
19.4414 | -104.7932
19.4451 | -104.7853
19.4488 | -104.7795
19.4525 | -104.7734
19.4562 | -104.7652
19.46 | -104.7574
19.4637 | -104.7541
19.4674 | -104.75
19.4711 | -104.7427
19.4748 | -104.7343
19.4785 | -104.7272
19.4822 | -104.7204
19.4859 | -104.7121
19.4897 | -104.7034
19.4934 | -104.6973
19.4971 | -104.6895
19.5008 | -104.6799
19.5045 | -104.6726
19.5082 | -104.6661
19.5119 | -104.663
19.5156 | -104.658
19.5193 | -104.653
19.5231 | -104.6487
19.5268 | -104.6456
19.5305 | -104.643
19.5342 | -104.6408
19.5379 | -104.64
19.5416 | -104.6399
19.5453 | -104.6377
19.549 | -104.6371
19.5528 | -104.6349
19.5565 | -104.6343
19.5602 | -104.6318
19.5639 | -104.6253
19.5676 | -104.6223
19.5713 | -104.6192
19.575 | -104.6165
19.5787 | -104.613
19.5824 | -104.6086
19.5862 | -104.6049
19.5899 | -104.5992
19.5936 | -104.5933
19.5973 | -104.5867
19.601 | -104.5813
19.6047 | -104.5752
19.6084 | -104.5681
19.6121 | -104.5611
19.6158 | -104.5548
19.6196 | -104.5467
19.6233 | -104.539
19.627 | -104.5294
19.6307 | -104.5207
19.6344 | -104.5109
19.6381 | -104.5006
19.6418 | -104.4891
19.6455 | -104.4761
19.6493 | -104.4655
19.653 | -104.4525
19.6567 | -104.4405
19.6604 | -104.4299
19.6641 | -104.4186
19.6678 | -104.4093
19.6715 | -104.3992
19.6752 | -104.3893
19.6789 | -104.3798
19.6827 | -104.37
19.6864 | -104.3623
19.6901 | -104.3568
19.6938 | -104.3524
19.6975 | -104.3487
19.7012 | -104.3467
19.7049 | -104.3457
19.7086 | -104.3463
19.7124 | -104.3455
19.7161 | -104.3474
19.7198 | -104.3504
19.7235 | -104.3542
19.7272 | -104.3577
19.7309 | -104.3613
19.7346 | -104.3672
19.7383 | -104.3742
19.742 | -104.3812
19.7458 | -104.3905
19.7495 | -104.4015
19.7532 | -104.4126
19.7569 | -104.4243
19.7606 | -104.4377
19.7643 | -104.4541
19.768 | -104.4688
19.7717 | -104.4835
19.7754 | -104.4996
19.7792 | -104.5132
19.7829 | -104.5254
19.7866 | -104.539
19.7903 | -104.553
19.794 | -104.5668
19.7977 | -104.5796
19.8014 | -104.5942
19.8051 | -104.6093
19.8089 | -104.6234
19.8126 | -104.6374
19.8163 | -104.6518
19.82 | -104.6651
19.8237 | -104.6779
19.8274 | -104.69
19.8311 | -104.7035
19.8348 | -104.7184
19.8385 | -104.7313
19.8423 | -104.7443
19.846 | -104.7596
19.8497 | -104.7722
19.8534 | -104.781
19.8571 | -104.7899
19.8608 | -104.802
19.8645 | -104.8113
19.8682 | -104.8161
19.8719 | -104.8203
19.8757 | -104.8237
19.8794 | -104.8287
19.8831 | -104.8299
19.8868 | -104.8298
19.8905 | -104.8332
19.8942 | -104.8342
19.8979 | -104.8324
19.9016 | -104.8308
19.9054 | -104.8305
19.9091 | -104.8283
19.9128 | -104.824
19.9165 | -104.8226
19.9202 | -104.8206
19.9239 | -104.8179
19.9276 | -104.8143
19.9313 | -104.8075
19.935 | -104.8031
19.9388 | -104.797
19.9425 | -104.7899
19.9462 | -104.7824
19.9499 | -104.7721
19.9536 | -104.7621
19.9573 | -104.7526
19.961 | -104.7432
19.9647 | -104.7349
19.9685 | -104.7239
19.9722 | -104.7113
19.9759 | -104.7
19.9796 | -104.6877
19.9833 | -104.6728
19.987 | -104.659
19.9907 | -104.6429
19.9944 | -104.6264
19.9981 | -104.6121
20.0019 | -104.5973
20.0056 | -104.5814
20.0093 | -104.5691
20.013 | -104.5578
20.0167 | -104.5438
20.0204 | -104.5292
20.0241 | -104.5145
20.0278 | -104.5002
20.0315 | -104.4873
20.0353 | -104.4723
20.039 | -104.4573
20.0427 | -104.444
20.0464 | -104.4297
20.0501 | -104.4171
20.0538 | -104.4041
20.0575 | -104.3906
20.0612 | -104.3808
20.065 | -104.3705
20.0687 | -104.3624
20.0724 | -104.3539
20.0761 | -104.3474
20.0798 | -104.341
20.0835 | -104.3323
20.0872 | -104.3236
20.0909 | -104.315
20.0946 | -104.3076
20.0984 | -104.3019
20.1021 | -104.2977
20.1058 | -104.2945
20.1095 | -104.2896
20.1132 | -104.2845
20.1169 | -104.2799
20.1206 | -104.2747
20.1243 | -104.2702
20.1281 | -104.2671
20.1318 | -104.2642
20.1355 | -104.2618
20.1392 | -104.2618
20.1429 | -104.2598
20.1466 | -104.2558
20.1503 | -104.2545
20.154 | -104.2551
20.1577 | -104.2536
20.1615 | -104.254
20.1652 | -104.2539
20.1689 | -104.2539
20.1726 | -104.2548
20.1763 | -104.2526
20.18 | -104.2516
20.1837 | -104.2516
20.1874 | -104.251
20.1911 | -104.2472
20.1949 | -104.2443
20.1986 | -104.2401
20.2023 | -104.2327
20.206 | -104.2244
20.2097 | -104.2171
20.2134 | -104.212
20.2171 | -104.2057
20.2208 | -104.1965
20.2246 | -104.1859
20.2283 | -104.1738
20.232 | -104.1632
20.2357 | -104.1537
20.2394 | -104.1428
20.2431 | -104.1285
20.2468 | -104.1141
20.2505 | -104.0995
20.2542 | -104.0844
20.258 | -104.0673
20.2617 | -104.0492
20.2654 | -104.0314
20.2691 | -104.015
20.2728 | -103.9971
20.2765 | -103.9783
20.2802 | -103.9579
20.2839 | -103.936
20.2876 | -103.9143
20.2914 | -103.8939
20.2951 | -103.8715
20.2988 | -103.8504
20.3025 | -103.8324
20.3062 | -103.8145
20.3099 | -103.7961
20.3136 | -103.7759
20.3173 | -103.7546
20.3211 | -103.7358
20.3248 | -103.7179
20.3285 | -103.6996
20.3322 | -103.6843
20.3359 | -103.6684
20.3396 | -103.6533
20.3433 | -103.639
20.347 | -103.6294
20.3507 | -103.6163
20.3545 | -103.6055
20.3582 | -103.5952
20.3619 | -103.5848
20.3656 | -103.5751
20.3693 | -103.5668
20.373 | -103.5615
20.3767 | -103.5561
20.3804 | -103.5526
20.3842 | -103.5502
20.3879 | -103.547
20.3916 | -103.5452
20.3953 | -103.5426
20.399 | -103.5403
20.4027 | -103.5393
20.4064 | -103.5397
20.4101 | -103.5418
20.4138 | -103.5452
20.4176 | -103.548
20.4213 | -103.5501
20.425 | -103.5515
20.4287 | -103.5524
20.4324 | -103.5556
20.4361 | -103.5585
20.4398 | -103.5587
20.4435 | -103.5611
20.4472 | -103.5646
20.451 | -103.5659
20.4547 | -103.565
20.4584 | -103.5627
20.4621 | -103.5601
20.4658 | -103.5586
20.4695 | -103.5557
20.4732 | -103.5525
20.4769 | -103.5501
20.4807 | -103.5481
20.4844 | -103.5457
20.4881 | -103.5433
20.4918 | -103.54
20.4955 | -103.5358
20.4992 | -103.5321
20.5029 | -103.5283
20.5066 | -103.5246
20.5103 | -103.5214
20.5141 | -103.518
20.5178 | -103.5156
20.5215 | -103.511
20.5252 | -103.5051
20.5289 | -103.4988
20.5326 | -103.4942
20.5363 | -103.4894
20.54 | -103.4836
20.5438 | -103.478
20.5475 | -103.4744
20.5512 | -103.4706
20.5549 | -103.4683
20.5586 | -103.4656
20.5623 | -103.4621
20.566 | -103.459
20.5697 | -103.4576
20.5734 | -103.4555
20.5772 | -103.4533
20.5809 | -103.4514
20.5846 | -103.4483
20.5883 | -103.4451
20.592 | -103.4417
20.5957 | -103.4408
20.5994 | -103.4394
20.6031 | -103.4383
20.6068 | -103.4382
20.6106 | -103.4393
20.6143 | -103.4423
20.618 | -103.4457
20.6217 | -103.447
20.6254 | -103.4487
20.6291 | -103.4509
20.6328 | -103.4532
20.6365 | -103.4556
20.6403 | -103.4578
20.644 | -103.4585
20.6477 | -103.4587
20.6514 | -103.4573
20.6551 | -103.4556
20.6588 | -103.4535
20.6625 | -103.4535
20.6662 | -103.4526
20.6699 | -103.4511
20.6737 | -103.4491
20.6774 | -103.4464
20.6811 | -103.4423
20.6848 | -103.4373
20.6885 | -103.4323
20.6922 | -103.4291
20.6959 | -103.4245
20.6996 | -103.4195
20.7034 | -103.4128
20.7071 | -103.4065
20.7108 | -103.4002
20.7145 | -103.3948
20.7182 | -103.3903
20.7219 | -103.3847
20.7256 | -103.3792
20.7293 | -103.3748
20.733 | -103.3714
20.7368 | -103.3661
20.7405 | -103.3616
20.7442 | -103.3548
20.7479 | -103.3499
20.7516 | -103.3472
20.7553 | -103.3424
20.759 | -103.3373
20.7627 | -103.3315
20.7664 | -103.3266
20.7702 | -103.319
20.7739 | -103.309
20.7776 | -103.2996
20.7813 | -103.2927
20.785 | -103.2864
20.7887 | -103.2786
20.7924 | -103.2699
20.7961 | -103.2615
20.7999 | -103.2544
20.8036 | -103.2475
20.8073 | -103.2388
20.811 | -103.2303
20.8147 | -103.2224
20.8184 | -103.2145
20.8221 | -103.2078
20.8258 | -103.2016
20.8295 | -103.1935
20.8333 | -103.1877
20.837 | -103.1826
20.8407 | -103.1761
20.8444 | -103.1681
20.8481 | -103.1594
20.8518 | -103.1527
20.8555 | -103.1474
20.8592 | -103.1418
20.8629 | -103.1347
20.8667 | -103.1261
20.8704 | -103.118
20.8741 | -103.1117
20.8778 | -103.1063
20.8815 | -103.099
20.8852 | -103.0942
20.8889 | -103.0875
20.8926 | -103.0802
20.8964 | -103.0759
20.9001 | -103.0737
20.9038 | -103.071
20.9075 | -103.0657
20.9112 | -103.0594
20.9149 | -103.0528
20.9186 | -103.049
20.9223 | -103.0464
20.926 | -103.0445
20.9298 | -103.044
20.9335 | -103.0447
20.9372 | -103.0452
20.9409 | -103.0454
20.9446 | -103.0464
20.9483 | -103.0479
20.952 | -103.05
20.9557 | -103.0545
20.9595 | -103.0574
20.9632 | -103.0611
20.9669 | -103.0666
20.9706 | -103.0734
20.9743 | -103.0814
20.978 | -103.0888
20.9817 | -103.0962
20.9854 | -103.1048
20.9891 | -103.1119
20.9929 | -103.1192
20.9966 | -103.1254
21.0003 | -103.1335
21.004 | -103.1413
21.0077 | -103.1482
21.0114 | -103.1536
21.0151 | -103.1597
21.0188 | -103.1681
21.0225 | -103.1765
21.0263 | -103.183
21.03 | -103.1894
21.0337 | -103.1968
21.0374 | -103.202
21.0411 | -103.2094
21.0448 | -103.2154
21.0485 | -103.2212
21.0522 | -103.2285
21.056 | -103.2348
21.0597 | -103.2399
21.0634 | -103.242
21.0671 | -103.248
21.0708 | -103.2544
21.0745 | -103.2577
21.0782 | -103.2594
21.0819 | -103.2618
21.0856 | -103.2647
21.0894 | -103.2653
21.0931 | -103.2649
21.0968 | -103.2649
21.1005 | -103.2635
21.1042 | -103.264
21.1079 | -103.2643
21.1116 | -103.2645
21.1153 | -103.2672
21.1191 | -103.2694
21.1228 | -103.2701
21.1265 | -103.272
21.1302 | -103.2699
21.1339 | -103.2692
21.1376 | -103.2695
21.1413 | -103.2714
21.145 | -103.2734
21.1487 | -103.2743
21.1525 | -103.2769
21.1562 | -103.2795
21.1599 | -103.2809
21.1636 | -103.2856
21.1673 | -103.2892
21.171 | -103.2919
21.1747 | -103.2942
21.1784 | -103.2967
21.1821 | -103.2996
21.1859 | -103.2993
21.1896 | -103.3014
21.1933 | -103.3031
21.197 | -103.303
21.2007 | -103.3033
21.2044 | -103.3044
21.2081 | -103.3066
21.2118 | -103.3078
21.2156 | -103.3088
21.2193 | -103.3095
21.223 | -103.3126
21.2267 | -103.3136
21.2304 | -103.3144
21.2341 | -103.3161
21.2378 | -103.3183
21.2415 | -103.3211
21.2452 | -103.3253
21.249 | -103.3303
21.2527 | -103.3354
21.2564 | -103.3408
21.2601 | -103.3471
21.2638 | -103.354
21.2675 | -103.3601
21.2712 | -103.366
21.2749 | -103.3742
21.2786 | -103.3844
21.2824 | -103.3955
21.2861 | -103.4067
21.2898 | -103.4204
21.2935 | -103.4345
21.2972 | -103.4499
21.3009 | -103.4655
21.3046 | -103.4822
21.3083 | -103.5011
21.3121 | -103.5208
21.3158 | -103.5403
21.3195 | -103.5583
21.3232 | -103.5745
21.3269 | -103.5938
21.3306 | -103.6147
21.3343 | -103.6345
21.338 | -103.6561
21.3417 | -103.677
21.3455 | -103.6977
21.3492 | -103.7173
21.3529 | -103.7359
21.3566 | -103.7537
21.3603 | -103.773
21.364 | -103.7941
21.3677 | -103.8122
21.3714 | -103.8303
21.3752 | -103.8495
21.3789 | -103.8684
21.3826 | -103.8858
21.3863 | -103.9017
21.39 | -103.9178
21.3937 | -103.9339
21.3974 | -103.9481
21.4011 | -103.9637
21.4048 | -103.9803
21.4086 | -103.993
21.4123 | -104.0049
21.416 | -104.0132
21.4197 | -104.023
21.4234 | -104.0323
21.4271 | -104.0428
21.4308 | -104.0541
21.4345 | -104.0663
21.4382 | -104.0768
21.442 | -104.0882
21.4457 | -104.0987
21.4494 | -104.1076
21.4531 | -104.1151
21.4568 | -104.1198
21.4605 | -104.127
21.4642 | -104.1364
21.4679 | -104.1442
21.4717 | -104.151
21.4754 | -104.1544
21.4791 | -104.1594
21.4828 | -104.165
21.4865 | -104.1687
21.4902 | -104.1766
21.4939 | -104.1822
21.4976 | -104.1845
21.5013 | -104.1871
21.5051 | -104.1892
21.5088 | -104.1919
21.5125 | -104.1933
21.5162 | -104.1944
21.5199 | -104.1962
21.5236 | -104.1959
21.5273 | -104.1926
21.531 | -104.1916
21.5347 | -104.1895
21.5385 | -104.1863
21.5422 | -104.1833
21.5459 | -104.1836
21.5496 | -104.1829
21.5533 | -104.1803
21.557 | -104.1791
21.5607 | -104.1762
21.5644 | -104.172
21.5682 | -104.1666
21.5719 | -104.1643
21.5756 | -104.162
21.5793 | -104.1578
21.583 | -104.1574
21.5867 | -104.1563
21.5904 | -104.1556
21.5941 | -104.152
21.5978 | -104.1478
21.6016 | -104.145
21.6053 | -104.1444
21.609 | -104.1409
21.6127 | -104.1372
21.6164 | -104.1349
21.6201 | -104.1324
21.6238 | -104.1312
21.6275 | -104.1317
21.6313 | -104.1297
21.635 | -104.1282
21.6387 | -104.1279
21.6424 | -104.1267
21.6461 | -104.1247
21.6498 | -104.124
21.6535 | -104.1238
21.6572 | -104.1229
21.6609 | -104.1207
21.6647 | -104.1187
21.6684 | -104.1153
21.6721 | -104.1132
21.6758 | -104.1111
21.6795 | -104.1081
21.6832 | -104.1042
21.6869 | -104.0998
21.6906 | -104.0966
21.6943 | -104.0943
21.6981 | -104.09
21.7018 | -104.0854
21.7055 | -104.0809
21.7092 | -104.0747
21.7129 | -104.0713
21.7166 | -104.0669
21.7203 | -104.0599
21.724 | -104.0545
21.7278 | -104.0485
21.7315 | -104.0439
21.7352 | -104.0387
21.7389 | -104.033
21.7426 | -104.0292
21.7463 | -104.0242
21.75 | -104.0171
21.7537 | -104.009
21.7574 | -104.0039
21.7612 | -103.9988
21.7649 | -103.9922
21.7686 | -103.9853
21.7723 | -103.9786
21.776 | -103.9707
21.7797 | -103.965
21.7834 | -103.9607
21.7871 | -103.9565
21.7909 | -103.9523
21.7946 | -103.9469
21.7983 | -103.9412
21.802 | -103.9367
21.8057 | -103.9337
21.8094 | -103.9288
21.8131 | -103.923
21.8168 | -103.9169
21.8205 | -103.912
21.8243 | -103.9088
21.828 | -103.9048
21.8317 | -103.9026
21.8354 | -103.9003
21.8391 | -103.8983
21.8428 | -103.8965
21.8465 | -103.8951
21.8502 | -103.8918
21.8539 | -103.8869
21.8577 | -103.8825
21.8614 | -103.8807
21.8651 | -103.8772
21.8688 | -103.8733
21.8725 | -103.8705
21.8762 | -103.869
21.8799 | -103.8666
21.8836 | -103.8645
21.8874 | -103.8635
21.8911 | -103.8596
21.8948 | -103.8562
21.8985 | -103.8548
21.9022 | -103.8549
21.9059 | -103.8558
21.9096 | -103.8535
21.9133 | -103.8506
21.917 | -103.8512
21.9208 | -103.8521
21.9245 | -103.8508
21.9282 | -103.8508
21.9319 | -103.8511
21.9356 | -103.8505
21.9393 | -103.851
21.943 | -103.8495
21.9467 | -103.8469
21.9505 | -103.8466
21.9542 | -103.8435
21.9579 | -103.8426
21.9616 | -103.8414
21.9653 | -103.8395
21.969 | -103.8392
21.9727 | -103.8413
21.9764 | -103.8418
21.9801 | -103.8421
21.9839 | -103.8404
21.9876 | -103.8371
21.9913 | -103.8356
21.995 | -103.8336
21.9987 | -103.8314
22.0024 | -103.8304
22.0061 | -103.8308
22.0098 | -103.83
22.0135 | -103.8297
22.0173 | -103.8292
22.021 | -103.8281
22.0247 | -103.8282
22.0284 | -103.8293
22.0321 | -103.8332
22.0358 | -103.8332
22.0395 | -103.8333
22.0432 | -103.8335
22.047 | -103.8333
22.0507 | -103.8321
22.0544 | -103.8311
22.0581 | -103.8304
22.0618 | -103.8332
22.0655 | -103.837
22.0692 | -103.8384
22.0729 | -103.8395
22.0766 | -103.8432
22.0804 | -103.8458
22.0841 | -103.8469
22.0878 | -103.848
22.0915 | -103.8479
22.0952 | -103.8479
22.0989 | -103.8492
22.1026 | -103.8519
22.1063 | -103.8539
22.11 | -103.8529
22.1138 | -103.8532
22.1175 | -103.8558
22.1212 | -103.8611
22.1249 | -103.8654
22.1286 | -103.8691
22.1323 | -103.8714
22.136 | -103.8729
22.1397 | -103.8781
22.1435 | -103.8799
22.1472 | -103.8831
22.1509 | -103.8851
22.1546 | -103.885
22.1583 | -103.8866
22.162 | -103.8907
22.1657 | -103.8965
22.1694 | -103.9023
22.1731 | -103.9046
22.1769 | -103.9081
22.1806 | -103.9129
22.1843 | -103.9176
22.188 | -103.9207
22.1917 | -103.9229
22.1954 | -103.9245
22.1991 | -103.9259
22.2028 | -103.9275
22.2066 | -103.9303
22.2103 | -103.9323
22.214 | -103.932
22.2177 | -103.934
22.2214 | -103.9337
22.2251 | -103.9316
22.2288 | -103.9311
22.2325 | -103.9317
22.2362 | -103.931
22.24 | -103.9277
22.2437 | -103.9259
22.2474 | -103.9246
22.2511 | -103.9239
22.2548 | -103.9213
22.2585 | -103.9186
22.2622 | -103.9133
22.2659 | -103.91
22.2696 | -103.9078
22.2734 | -103.9029
22.2771 | -103.8976
22.2808 | -103.8931
22.2845 | -103.8879
22.2882 | -103.8822
22.2919 | -103.8787
22.2956 | -103.8758
22.2993 | -103.8707
22.3031 | -103.8687
22.3068 | -103.8653
22.3105 | -103.8617
22.3142 | -103.8594
22.3179 | -103.8558
22.3216 | -103.8484
22.3253 | -103.8446
22.329 | -103.8422
22.3327 | -103.8399
22.3365 | -103.8378
22.3402 | -103.8338
22.3439 | -103.831
22.3476 | -103.8287
22.3513 | -103.8281
22.355 | -103.8279
22.3587 | -103.8284
22.3624 | -103.83
22.3662 | -103.8326
22.3699 | -103.8336
22.3736 | -103.8341
22.3773 | -103.8343
22.381 | -103.835
22.3847 | -103.8338
22.3884 | -103.8315
22.3921 | -103.8241
22.3958 | -103.8127
22.3996 | -103.7983
22.4033 | -103.7834
22.407 | -103.7718
22.4107 | -103.7595
22.4144 | -103.7449
22.4181 | -103.7283
22.4218 | -103.7115
22.4255 | -103.6946
22.4292 | -103.6747
22.433 | -103.6504
22.4367 | -103.6243
22.4404 | -103.5989
22.4441 | -103.571
22.4478 | -103.543
22.4515 | -103.5107
22.4552 | -103.4732
22.4589 | -103.4359
22.4627 | -103.3986
22.4664 | -103.3594
22.4701 | -103.3183
22.4738 | -103.2741
22.4775 | -103.2287
22.4812 | -103.1802
22.4849 | -103.1317
22.4886 | -103.0732
22.4923 | -103.0255
22.4961 | -102.9793
22.4998 | -102.9315
22.5035 | -102.8831
22.5072 | -102.8367
22.5109 | -102.7906
22.5146 | -102.7442
22.5183 | -102.6985
22.522 | -102.6541
22.5257 | -102.6103
22.5295 | -102.5675
22.5332 | -102.5258
22.5369 | -102.4859
22.5406 | -102.4506
22.5443 | -102.4181
22.548 | -102.3859
22.5517 | -102.3552
22.5554 | -102.3281
22.5592 | -102.3042
22.5629 | -102.2822
22.5666 | -102.2632
22.5703 | -102.2462
22.574 | -102.2343
22.5777 | -102.2276
22.5814 | -102.2233
22.5851 | -102.2201
22.5888 | -102.2203
22.5926 | -102.2248
22.5963 | -102.2299
22.6 | -102.238
22.6037 | -102.2495
22.6074 | -102.2626
22.6111 | -102.2762
22.6148 | -102.2922
22.6185 | -102.312
22.6223 | -102.3333
22.626 | -102.3566
22.6297 | -102.3821
22.6334 | -102.4076
22.6371 | -102.4351
22.6408 | -102.4652
22.6445 | -102.4945
22.6482 | -102.5241
22.6519 | -102.5544
22.6557 | -102.5838
22.6594 | -102.6144
22.6631 | -102.6448
22.6668 | -102.6741
22.6705 | -102.7028
22.6742 | -102.7327
22.6779 | -102.76
22.6816 | -102.7854
22.6853 | -102.8095
22.6891 | -102.8315
22.6928 | -102.8529
22.6965 | -102.8739
22.7002 | -102.8948
22.7039 | -102.9154
22.7076 | -102.9359
22.7113 | -102.9577
22.715 | -102.9791
22.7188 | -102.9997
22.7225 | -103.0176
22.7262 | -103.0303
22.7299 | -103.0384
22.7336 | -103.0436
22.7373 | -103.0492
22.741 | -103.0541
22.7447 | -103.0581
22.7484 | -103.0602
22.7522 | -103.06
22.7559 | -103.0573
22.7596 | -103.0536
22.7633 | -103.0501
22.767 | -103.0464
22.7707 | -103.0451
22.7744 | -103.0435
22.7781 | -103.0396
22.7818 | -103.031
22.7856 | -103.024
22.7893 | -103.0203
22.793 | -103.0147
22.7967 | -103.007
22.8004 | -103.0007
22.8041 | -102.9929
22.8078 | -102.9858
22.8115 | -102.9816
22.8153 | -102.9763
22.819 | -102.9692
22.8227 | -102.9627
22.8264 | -102.9579
22.8301 | -102.9515
22.8338 | -102.9455
22.8375 | -102.9414
22.8412 | -102.9374
22.8449 | -102.9317
22.8487 | -102.924
22.8524 | -102.9177
22.8561 | -102.9128
22.8598 | -102.9066
22.8635 | -102.8994
22.8672 | -102.8918
22.8709 | -102.8856
22.8746 | -102.8791
22.8784 | -102.8738
22.8821 | -102.867
22.8858 | -102.8592
22.8895 | -102.8503
22.8932 | -102.8405
22.8969 | -102.8276
22.9006 | -102.8133
22.9043 | -102.8008
22.908 | -102.79
22.9118 | -102.7774
22.9155 | -102.7633
22.9192 | -102.751
22.9229 | -102.7398
22.9266 | -102.7258
22.9303 | -102.7124
22.934 | -102.6995
22.9377 | -102.6872
22.9414 | -102.6735
22.9452 | -102.657
22.9489 | -102.6398
22.9526 | -102.6223
22.9563 | -102.6072
22.96 | -102.5924
22.9637 | -102.5767
22.9674 | -102.5604
22.9711 | -102.5422
22.9749 | -102.5237
22.9786 | -102.506
22.9823 | -102.4891
22.986 | -102.472
22.9897 | -102.4557
22.9934 | -102.4403
22.9971 | -102.4235
23.0008 | -102.409
23.0045 | -102.3968
23.0083 | -102.3847
23.012 | -102.3725
23.0157 | -102.3585
23.0194 | -102.3431
23.0231 | -102.3332
23.0268 | -102.3184
23.0305 | -102.306
23.0342 | -102.2928
23.038 | -102.2793
23.0417 | -102.2678
23.0454 | -102.2585
23.0491 | -102.2472
23.0528 | -102.2356
23.0565 | -102.226
23.0602 | -102.2185
23.0639 | -102.2101
23.0676 | -102.202
23.0714 | -102.1937
23.0751 | -102.1874
23.0788 | -102.1812
23.0825 | -102.1764
23.0862 | -102.1712
23.0899 | -102.1667
23.0936 | -102.1608
23.0973 | -102.1553
23.101 | -102.1508
23.1048 | -102.1458
23.1085 | -102.1424
23.1122 | -102.1382
23.1159 | -102.1304
23.1196 | -102.1248
23.1233 | -102.1199
23.127 | -102.1114
23.1307 | -102.105
23.1345 | -102.0992
23.1382 | -102.0928
23.1419 | -102.0865
23.1456 | -102.0794
23.1493 | -102.07
23.153 | -102.0624
23.1567 | -102.0543
23.1604 | -102.0459
23.1641 | -102.0369
23.1679 | -102.0277
23.1716 | -102.0188
23.1753 | -102.01
23.179 | -101.9986
23.1827 | -101.9865
23.1864 | -101.9753
23.1901 | -101.9644
23.1938 | -101.9514
23.1975 | -101.9353
23.2013 | -101.9216
23.205 | -101.9088
23.2087 | -101.8964
23.2124 | -101.8836
23.2161 | -101.8703
23.2198 | -101.8578
23.2235 | -101.8448
23.2272 | -101.8304
23.231 | -101.8172
23.2347 | -101.8025
23.2384 | -101.7876
23.2421 | -101.7721
23.2458 | -101.7572
23.2495 | -101.739
23.2532 | -101.7243
23.2569 | -101.7099
23.2606 | -101.6953
23.2644 | -101.6803
23.2681 | -101.6659
23.2718 | -101.6517
23.2755 | -101.6378
23.2792 | -101.6248
23.2829 | -101.6117
23.2866 | -101.5987
23.2903 | -101.5867
23.2941 | -101.5753
23.2978 | -101.5652
23.3015 | -101.5559
23.3052 | -101.5467
23.3089 | -101.5374
23.3126 | -101.5286
23.3163 | -101.5201
23.32 | -101.513
23.3237 | -101.507
23.3275 | -101.4998
23.3312 | -101.4933
23.3349 | -101.4875
23.3386 | -101.483
23.3423 | -101.4802
23.346 | -101.4798
23.3497 | -101.478
23.3534 | -101.4764
23.3571 | -101.4782
23.3609 | -101.4794
23.3646 | -101.4822
23.3683 | -101.4867
23.372 | -101.4915
23.3757 | -101.4949
23.3794 | -101.5001
23.3831 | -101.5065
23.3868 | -101.5137
23.3906 | -101.5212
23.3943 | -101.5283
23.398 | -101.5333
23.4017 | -101.5395
23.4054 | -101.5466
23.4091 | -101.556
23.4128 | -101.5636
23.4165 | -101.5709
23.4202 | -101.5783
23.424 | -101.5857
23.4277 | -101.5902
23.4314 | -101.594
23.4351 | -101.5994
23.4388 | -101.6054
23.4425 | -101.6102
23.4462 | -101.6168
23.4499 | -101.6197
23.4537 | -101.6225
23.4574 | -101.6271
23.4611 | -101.6304
23.4648 | -101.6327
23.4685 | -101.6341
23.4722 | -101.6354
23.4759 | -101.6379
23.4796 | -101.6423
23.4833 | -101.6417
23.4871 | -101.6414
23.4908 | -101.6413
23.4945 | -101.6407
23.4982 | -101.6413
23.5019 | -101.6429
23.5056 | -101.6435
23.5093 | -101.6435
23.513 | -101.644
23.5167 | -101.642
23.5205 | -101.6381
23.5242 | -101.6355
23.5279 | -101.6352
23.5316 | -101.6323
23.5353 | -101.6285
23.539 | -101.6243
23.5427 | -101.6168
23.5464 | -101.609
23.5502 | -101.606
23.5539 | -101.6035
23.5576 | -101.6016
23.5613 | -101.5997
23.565 | -101.5952
23.5687 | -101.5912
23.5724 | -101.5871
23.5761 | -101.5837
23.5798 | -101.5802
23.5836 | -101.5792
23.5873 | -101.5764
23.591 | -101.574
23.5947 | -101.5716
23.5984 | -101.5699
23.6021 | -101.5686
23.6058 | -101.568
23.6095 | -101.5654
23.6133 | -101.5638
23.617 | -101.5628
23.6207 | -101.56
23.6244 | -101.5579
23.6281 | -101.5563
23.6318 | -101.5549
23.6355 | -101.5559
23.6392 | -101.5551
23.6429 | -101.5538
23.6467 | -101.5537
23.6504 | -101.5546
23.6541 | -101.554
23.6578 | -101.5515
23.6615 | -101.5511
23.6652 | -101.5503
23.6689 | -101.5475
23.6726 | -101.5444
23.6763 | -101.5441
23.6801 | -101.541
23.6838 | -101.5388
23.6875 | -101.5336
23.6912 | -101.5234
23.6949 | -101.517
23.6986 | -101.5086
23.7023 | -101.5016
23.706 | -101.4952
23.7098 | -101.4879
23.7135 | -101.481
23.7172 | -101.4734
23.7209 | -101.4642
23.7246 | -101.4541
23.7283 | -101.4462
23.732 | -101.44
23.7357 | -101.4306
23.7394 | -101.4199
23.7432 | -101.407
23.7469 | -101.3957
23.7506 | -101.385
23.7543 | -101.3726
23.758 | -101.3616
23.7617 | -101.353
23.7654 | -101.3437
23.7691 | -101.3348
23.7728 | -101.3257
23.7766 | -101.3167
23.7803 | -101.3081
23.784 | -101.3009
23.7877 | -101.2939
23.7914 | -101.2862
23.7951 | -101.2786
23.7988 | -101.2703
23.8025 | -101.2625
23.8063 | -101.2553
23.81 | -101.2489
23.8137 | -101.2422
23.8174 | -101.238
23.8211 | -101.2349
23.8248 | -101.232
23.8285 | -101.2294
23.8322 | -101.2255
23.8359 | -101.2211
23.8397 | -101.2211
23.8434 | -101.2223
23.8471 | -101.2199
23.8508 | -101.2205
23.8545 | -101.2247
23.8582 | -101.227
23.8619 | -101.2305
23.8656 | -101.2327
23.8694 | -101.2327
23.8731 | -101.2361
23.8768 | -101.245
23.8805 | -101.2531
23.8842 | -101.2572
23.8879 | -101.2631
23.8916 | -101.2695
23.8953 | -101.2764
23.899 | -101.2831
23.9028 | -101.2893
23.9065 | -101.2951
23.9102 | -101.3015
23.9139 | -101.3081
23.9176 | -101.3153
23.9213 | -101.3197
23.925 | -101.3217
23.9287 | -101.3231
23.9324 | -101.3263
23.9362 | -101.3294
23.9399 | -101.3324
23.9436 | -101.3356
23.9473 | -101.3395
23.951 | -101.3406
23.9547 | -101.3391
23.9584 | -101.34
23.9621 | -101.3407
23.9659 | -101.3417
23.9696 | -101.3407
23.9733 | -101.3375
23.977 | -101.3356
23.9807 | -101.3334
23.9844 | -101.3289
23.9881 | -101.3236
23.9918 | -101.3138
23.9955 | -101.3038
23.9993 | -101.2985
24.003 | -101.2919
24.0067 | -101.2851
24.0104 | -101.2809
24.0141 | -101.2726
24.0178 | -101.2641
24.0215 | -101.2565
24.0252 | -101.2518
24.029 | -101.2464
24.0327 | -101.2388
24.0364 | -101.2327
24.0401 | -101.2239
24.0438 | -101.2143
24.0475 | -101.2058
24.0512 | -101.1945
24.0549 | -101.1841
24.0586 | -101.1748
24.0624 | -101.1638
24.0661 | -101.1536
24.0698 | -101.1455
24.0735 | -101.1386
24.0772 | -101.1314
24.0809 | -101.1266
24.0846 | -101.1206
24.0883 | -101.1127
24.092 | -101.1066
24.0958 | -101.1027
24.0995 | -101.0946
24.1032 | -101.0864
24.1069 | -101.0772
24.1106 | -101.0675
24.1143 | -101.0562
24.118 | -101.0463
24.1217 | -101.037
24.1255 | -101.0285
24.1292 | -101.0197
24.1329 | -101.0126
24.1366 | -101.0036
24.1403 | -100.9964
24.144 | -100.9919
24.1477 | -100.9874
24.1514 | -100.9829
24.1551 | -100.9776
24.1589 | -100.9724
24.1626 | -100.9673
24.1663 | -100.9625
24.17 | -100.9572
24.1737 | -100.952
24.1774 | -100.9499
24.1811 | -100.9452
24.1848 | -100.942
24.1885 | -100.9424
24.1923 | -100.9391
24.196 | -100.9335
24.1997 | -100.9316
24.2034 | -100.9267
24.2071 | -100.9218
24.2108 | -100.9221
24.2145 | -100.9191
24.2182 | -100.9141
24.222 | -100.9114
24.2257 | -100.9098
24.2294 | -100.9052
24.2331 | -100.9012
24.2368 | -100.8967
24.2405 | -100.8923
24.2442 | -100.89
24.2479 | -100.8885
24.2516 | -100.8823
24.2554 | -100.8732
24.2591 | -100.8595
24.2628 | -100.8458
24.2665 | -100.8265
24.2702 | -100.8069
24.2739 | -100.7686
24.2776 | -100.7147
24.2813 | -100.6692
24.2851 | -100.6188
24.2888 | -100.5699
24.2925 | -100.5233
24.2962 | -100.4752
24.2999 | -100.4265
24.3036 | -100.3742
24.3073 | -100.3162
24.311 | -100.2579
24.3147 | -100.1931
24.3185 | -100.128
24.3222 | -100.0648
24.3259 | -99.9998
24.3296 | -99.9311
24.3333 | -99.8639
24.337 | -99.7953
24.3407 | -99.7186
24.3444 | -99.6414
24.3481 | -99.5633
24.3519 | -99.4802
24.3556 | -99.3974
24.3593 | -99.3148
24.363 | -99.2303
24.3667 | -99.1468
24.3704 | -99.0646
24.3741 | -98.9821
24.3778 | -98.9013
24.3816 | -98.8229
24.3853 | -98.744
24.389 | -98.6668
24.3927 | -98.5948
24.3964 | -98.5237
24.4001 | -98.4555
24.4038 | -98.3916
24.4075 | -98.3291
24.4112 | -98.2694
24.415 | -98.2135
24.4187 | -98.1608
24.4224 | -98.1127
24.4261 | -98.0697
24.4298 | -98.0319
24.4335 | -97.9985
24.4372 | -97.969
24.4409 | -97.9447
24.4447 | -97.9261
24.4484 | -97.9129
24.4521 | -97.9049
24.4558 | -97.9012
24.4595 | -97.9039
24.4632 | -97.911
24.4669 | -97.9223
24.4706 | -97.9379
24.4743 | -97.9587
24.4781 | -97.9838
24.4818 | -98.0146
24.4855 | -98.0485
24.4892 | -98.0863
24.4929 | -98.1293
24.4966 | -98.1766
24.5003 | -98.2289
24.504 | -98.285
24.5077 | -98.3435
24.5115 | -98.4054
24.5152 | -98.4707
24.5189 | -98.5381
24.5226 | -98.6063
24.5263 | -98.6775
24.53 | -98.7494
24.5337 | -98.8223
24.5374 | -98.8989
24.5412 | -98.9775
24.5449 | -99.0541
24.5486 | -99.1316
24.5523 | -99.2101
24.556 | -99.2857
24.5597 | -99.3613
24.5634 | -99.4347
24.5671 | -99.5013
24.5708 | -99.5544
24.5746 | -99.608
24.5783 | -99.6563
24.582 | -99.7051
24.5857 | -99.7448
24.5894 | -99.7616
24.5931 | -99.7535
24.5968 | -99.7446
24.6005 | -99.7405
24.6043 | -99.7391
24.608 | -99.7339
24.6117 | -99.7233
24.6154 | -99.704
24.6191 | -99.6827
24.6228 | -99.6629
24.6265 | -99.6295
24.6302 | -99.5216
24.6339 | -99.4749
24.6377 | -99.4179
24.6414 | -99.3503
24.6451 | -99.2766
24.6488 | -99.1974
24.6525 | -99.1143
24.6562 | -99.0294
24.6599 | -98.9402
24.6636 | -98.8499
24.6673 | -98.7563
24.6711 | -98.6585
24.6748 | -98.5578
24.6785 | -98.4559
24.6822 | -98.3564
24.6859 | -98.2574
24.6896 | -98.1612
24.6933 | -98.0674
24.697 | -97.974
24.7008 | -97.8848
24.7045 | -97.7984
24.7082 | -97.7158
24.7119 | -97.6373
24.7156 | -97.5623
24.7193 | -97.4908
24.723 | -97.4226
24.7267 | -97.36
24.7304 | -97.3019
24.7342 | -97.2496
24.7379 | -97.2032
24.7416 | -97.161
24.7453 | -97.1238
24.749 | -97.0922
24.7527 | -97.0656
24.7564 | -97.046
24.7601 | -97.0323
24.7638 | -97.0243
24.7676 | -97.0227
24.7713 | -97.026
24.775 | -97.0339
24.7787 | -97.0475
24.7824 | -97.0681
24.7861 | -97.0946
24.7898 | -97.1268
24.7935 | -97.1644
24.7973 | -97.207
24.801 | -97.2544
24.8047 | -97.3078
24.8084 | -97.3658
24.8121 | -97.4283
24.8158 | -97.4951
24.8195 | -97.5664
24.8232 | -97.6414
24.8269 | -97.7206
24.8307 | -97.8028
24.8344 | -97.8896
24.8381 | -97.9792
24.8418 | -98.0699
24.8455 | -98.1591
24.8492 | -98.252
24.8529 | -98.3476
24.8566 | -98.4464
24.8604 | -98.5462
24.8641 | -98.6472
24.8678 | -98.7499
24.8715 | -98.8533
24.8752 | -98.9555
24.8789 | -99.0575
24.8826 | -99.1591
24.8863 | -99.2579
24.89 | -99.3537
24.8938 | -99.4488
24.8975 | -99.5406
24.9012 | -99.6282
24.9049 | -99.7125
24.9086 | -99.7936
24.9123 | -99.8718
24.916 | -99.9533
24.9197 | -100.0409
24.9234 | -100.1209
24.9272 | -100.1958
24.9309 | -100.2616
24.9346 | -100.327
24.9383 | -100.4025
24.942 | -100.4901
24.9457 | -100.5626
24.9494 | -100.6125
24.9531 | -100.6464
24.9569 | -100.6719
24.9606 | -100.6958
24.9643 | -100.7151
24.968 | -100.7364
24.9717 | -100.7492
24.9754 | -100.7516
24.9791 | -100.7517
24.9828 | -100.7522
24.9865 | -100.7517
24.9903 | -100.748
24.994 | -100.7453
24.9977 | -100.7432
25.0014 | -100.7385
25.0051 | -100.7339
25.0088 | -100.7289
25.0125 | -100.723
25.0162 | -100.7191
25.02 | -100.7137
25.0237 | -100.7069
25.0274 | -100.7046
25.0311 | -100.7041
25.0348 | -100.6987
25.0385 | -100.6942
25.0422 | -100.6896
25.0459 | -100.6838
25.0496 | -100.6818
25.0534 | -100.6788
25.0571 | -100.6743
25.0608 | -100.6698
25.0645 | -100.6661
25.0682 | -100.6629
25.0719 | -100.6622
25.0756 | -100.6641
25.0793 | -100.6645
25.083 | -100.6654
25.0868 | -100.6696
25.0905 | -100.6734
25.0942 | -100.6783
25.0979 | -100.6865
25.1016 | -100.6952
25.1053 | -100.7053
25.109 | -100.714
25.1127 | -100.7238
25.1165 | -100.7364
25.1202 | -100.7494
25.1239 | -100.7629
25.1276 | -100.7789
25.1313 | -100.7955
25.135 | -100.8093
25.1387 | -100.824
25.1424 | -100.8394
25.1461 | -100.8551
25.1499 | -100.8707
25.1536 | -100.8895
25.1573 | -100.9086
25.161 | -100.9268
25.1647 | -100.9481
25.1684 | -100.9651
25.1721 | -100.9852
25.1758 | -101.0066
25.1795 | -101.0283
25.1833 | -101.0496
25.187 | -101.071
25.1907 | -101.0903
25.1944 | -101.1083
25.1981 | -101.1281
25.2018 | -101.1549
25.2055 | -101.1765
25.2092 | -101.1968
25.213 | -101.2166
25.2167 | -101.236
25.2204 | -101.255
25.2241 | -101.2733
25.2278 | -101.2879
25.2315 | -101.3033
25.2352 | -101.3204
25.2389 | -101.336
25.2426 | -101.3517
25.2464 | -101.3678
25.2501 | -101.3824
25.2538 | -101.4001
25.2575 | -101.4199
25.2612 | -101.4396
25.2649 | -101.4569
25.2686 | -101.4743
25.2723 | -101.4876
25.2761 | -101.4993
25.2798 | -101.5086
25.2835 | -101.5198
25.2872 | -101.5294
25.2909 | -101.5399
25.2946 | -101.5512
25.2983 | -101.5604
25.302 | -101.5688
25.3057 | -101.5776
25.3095 | -101.584
25.3132 | -101.5898
25.3169 | -101.5946
25.3206 | -101.5993
25.3243 | -101.6037
25.328 | -101.6069
25.3317 | -101.6113
25.3354 | -101.6138
25.3391 | -101.6153
25.3429 | -101.6168
25.3466 | -101.6182
25.3503 | -101.6181
25.354 | -101.6173
25.3577 | -101.6148
25.3614 | -101.6122
25.3651 | -101.6103
25.3688 | -101.6095
25.3726 | -101.6066
25.3763 | -101.6036
25.38 | -101.6026
25.3837 | -101.6001
25.3874 | -101.599
25.3911 | -101.5976
25.3948 | -101.5947
25.3985 | -101.5907
25.4022 | -101.5854
25.406 | -101.5813
25.4097 | -101.5768
25.4134 | -101.5739
25.4171 | -101.57
25.4208 | -101.5652
25.4245 | -101.5596
25.4282 | -101.5547
25.4319 | -101.5483
25.4356 | -101.5414
25.4394 | -101.5346
25.4431 | -101.5274
25.4468 | -101.5226
25.4505 | -101.5159
25.4542 | -101.5086
25.4579 | -101.5018
25.4616 | -101.496
25.4653 | -101.4916
25.4691 | -101.4865
25.4728 | -101.4806
25.4765 | -101.4769
25.4802 | -101.4751
25.4839 | -101.472
25.4876 | -101.4674
25.4913 | -101.4621
25.495 | -101.4581
25.4987 | -101.456
25.5025 | -101.4515
25.5062 | -101.4488
25.5099 | -101.4455
25.5136 | -101.4398
25.5173 | -101.4361
25.521 | -101.4334
25.5247 | -101.4312
25.5284 | -101.4324
25.5322 | -101.4324
25.5359 | -101.4314
25.5396 | -101.4307
25.5433 | -101.4294
25.547 | -101.428
25.5507 | -101.429
25.5544 | -101.4306
25.5581 | -101.4311
25.5618 | -101.4319
25.5656 | -101.4334
25.5693 | -101.4328
25.573 | -101.4315
25.5767 | -101.4317
25.5804 | -101.4317
25.5841 | -101.4318
25.5878 | -101.4331
25.5915 | -101.4334
25.5952 | -101.4343
25.599 | -101.4377
25.6027 | -101.44
25.6064 | -101.4415
25.6101 | -101.4442
25.6138 | -101.4461
25.6175 | -101.4463
25.6212 | -101.4469
25.6249 | -101.4475
25.6287 | -101.4478
25.6324 | -101.4474
25.6361 | -101.4479
25.6398 | -101.4497
25.6435 | -101.4496
25.6472 | -101.4508
25.6509 | -101.4508
25.6546 | -101.4527
25.6583 | -101.4522
25.6621 | -101.453
25.6658 | -101.454
25.6695 | -101.4538
25.6732 | -101.4547
25.6769 | -101.4556
25.6806 | -101.4533
25.6843 | -101.4505
25.688 | -101.451
25.6918 | -101.4497
25.6955 | -101.4473
25.6992 | -101.4408
25.7029 | -101.4361
25.7066 | -101.4342
25.7103 | -101.4332
25.714 | -101.4308
25.7177 | -101.4252
25.7214 | -101.4206
25.7252 | -101.4149
25.7289 | -101.4087
25.7326 | -101.3986
25.7363 | -101.3877
25.74 | -101.3814
25.7437 | -101.3721
25.7474 | -101.3622
25.7511 | -101.3534
25.7548 | -101.3439
25.7586 | -101.3353
25.7623 | -101.3283
25.766 | -101.3188
25.7697 | -101.3076
25.7734 | -101.2986
25.7771 | -101.2877
25.7808 | -101.2759
25.7845 | -101.2647
25.7883 | -101.2522
25.792 | -101.2379
25.7957 | -101.2238
25.7994 | -101.2108
25.8031 | -101.1974
25.8068 | -101.1818
25.8105 | -101.1646
25.8142 | -101.1512
25.8179 | -101.1374
25.8217 | -101.1213
25.8254 | -101.1024
25.8291 | -101.086
25.8328 | -101.068
25.8365 | -101.0469
25.8402 | -101.0285
25.8439 | -101.0102
25.8476 | -100.9918
25.8514 | -100.9759
25.8551 | -100.961
25.8588 | -100.9447
25.8625 | -100.9277
25.8662 | -100.9096
25.8699 | -100.8888
25.8736 | -100.8715
25.8773 | -100.8549
25.881 | -100.8367
25.8848 | -100.8191
25.8885 | -100.8031
25.8922 | -100.7858
25.8959 | -100.7683
25.8996 | -100.7508
25.9033 | -100.7324
25.907 | -100.7161
25.9107 | -100.702
25.9144 | -100.6875
25.9182 | -100.6735
25.9219 | -100.6606
25.9256 | -100.6468
25.9293 | -100.6331
25.933 | -100.6198
25.9367 | -100.6077
25.9404 | -100.5945
25.9441 | -100.5814
25.9479 | -100.5693
25.9516 | -100.5601
25.9553 | -100.5522
25.959 | -100.5434
25.9627 | -100.5343
25.9664 | -100.5246
25.9701 | -100.5153
25.9738 | -100.5069
25.9775 | -100.4993
25.9813 | -100.4922
25.985 | -100.4847
25.9887 | -100.4784
25.9924 | -100.4713
25.9961 | -100.4646
25.9998 | -100.4585
26.0035 | -100.4527
26.0072 | -100.448
26.0109 | -100.4447
26.0147 | -100.4393
26.0184 | -100.4339
26.0221 | -100.432
26.0258 | -100.4291
26.0295 | -100.4254
26.0332 | -100.4244
26.0369 | -100.4226
26.0406 | -100.4211
26.0444 | -100.4217
26.0481 | -100.4203
26.0518 | -100.4171
26.0555 | -100.4167
26.0592 | -100.4171
26.0629 | -100.4174
26.0666 | -100.4199
26.0703 | -100.4197
26.074 | -100.4164
26.0778 | -100.4151
26.0815 | -100.4159
26.0852 | -100.4175
26.0889 | -100.4186
26.0926 | -100.4196
26.0963 | -100.4187
26.1 | -100.4204
26.1037 | -100.4237
26.1075 | -100.422
26.1112 | -100.4214
26.1149 | -100.4207
26.1186 | -100.4192
26.1223 | -100.4193
26.126 | -100.4193
26.1297 | -100.4178
26.1334 | -100.418
26.1371 | -100.4183
26.1409 | -100.4156
26.1446 | -100.4152
26.1483 | -100.4155
26.152 | -100.4152
26.1557 | -100.4163
26.1594 | -100.4177
26.1631 | -100.4192
26.1668 | -100.4191
26.1705 | -100.4194
26.1743 | -100.4191
26.178 | -100.4172
26.1817 | -100.4144
26.1854 | -100.4122
26.1891 | -100.412
26.1928 | -100.4136
26.1965 | -100.4126
26.2002 | -100.4132
26.204 | -100.4154
26.2077 | -100.4114
26.2114 | -100.4075
26.2151 | -100.4058
26.2188 | -100.4039
26.2225 | -100.4021
26.2262 | -100.4013
26.2299 | -100.3996
26.2336 | -100.3996
26.2374 | -100.3991
26.2411 | -100.3998
26.2448 | -100.3988
26.2485 | -100.3973
26.2522 | -100.3957
26.2559 | -100.3937
26.2596 | -100.3926
26.2633 | -100.3913
26.2671 | -100.3913
26.2708 | -100.3923
26.2745 | -100.3951
26.2782 | -100.3964
26.2819 | -100.3953
26.2856 | -100.3928
26.2893 | -100.3895
26.293 | -100.3877
26.2967 | -100.3862
26.3005 | -100.3857
26.3042 | -100.3868
26.3079 | -100.3868
26.3116 | -100.3881
26.3153 | -100.3901
26.319 | -100.3916
26.3227 | -100.3921
26.3264 | -100.391
26.3301 | -100.3906
26.3339 | -100.391
26.3376 | -100.3929
26.3413 | -100.3942
26.345 | -100.3957
26.3487 | -100.396
26.3524 | -100.3956
26.3561 | -100.396
26.3598 | -100.3975
26.3636 | -100.3992
26.3673 | -100.4004
26.371 | -100.4017
26.3747 | -100.404
26.3784 | -100.406
26.3821 | -100.4071
26.3858 | -100.4096
26.3895 | -100.4119
26.3932 | -100.4131
26.397 | -100.4128
26.4007 | -100.4122
26.4044 | -100.4128
26.4081 | -100.4137
26.4118 | -100.4157
26.4155 | -100.4171
26.4192 | -100.4208
26.4229 | -100.4245
26.4266 | -100.4275
26.4304 | -100.4308
26.4341 | -100.4342
26.4378 | -100.4381
26.4415 | -100.4443
26.4452 | -100.4499
26.4489 | -100.4534
26.4526 | -100.4552
26.4563 | -100.4558
26.4601 | -100.4576
26.4638 | -100.46
26.4675 | -100.4646
26.4712 | -100.4686
26.4749 | -100.472
26.4786 | -100.4751
26.4823 | -100.4777
26.486 | -100.4829
26.4897 | -100.487
26.4935 | -100.4889
26.4972 | -100.4897
26.5009 | -100.4914
26.5046 | -100.494
26.5083 | -100.4974
26.512 | -100.4999
26.5157 | -100.5014
26.5194 | -100.5033
26.5232 | -100.5058
26.5269 | -100.5086
26.5306 | -100.5119
26.5343 | -100.5155
26.538 | -100.5192
26.5417 | -100.5216
26.5454 | -100.5255
26.5491 | -100.5307
26.5528 | -100.5348
26.5566 | -100.5385
26.5603 | -100.5407
26.564 | -100.5422
26.5677 | -100.5434
26.5714 | -100.5458
26.5751 | -100.5494
26.5788 | -100.5547
26.5825 | -100.559
26.5862 | -100.5637
26.59 | -100.5674
26.5937 | -100.5703
26.5974 | -100.5728
26.6011 | -100.5768
26.6048 | -100.5827
26.6085 | -100.5877
26.6122 | -100.593
26.6159 | -100.5979
26.6197 | -100.6023
26.6234 | -100.6086
26.6271 | -100.6158
26.6308 | -100.6226
26.6345 | -100.6286
26.6382 | -100.6364
26.6419 | -100.6439
26.6456 | -100.6499
26.6493 | -100.6556
26.6531 | -100.6634
26.6568 | -100.671
26.6605 | -100.6804
26.6642 | -100.6932
26.6679 | -100.7048
26.6716 | -100.7164
26.6753 | -100.7277
26.679 | -100.7384
26.6827 | -100.749
26.6865 | -100.7603
26.6902 | -100.7702
26.6939 | -100.7787
26.6976 | -100.7896
26.7013 | -100.8029
26.705 | -100.814
26.7087 | -100.8242
26.7124 | -100.8333
26.7162 | -100.8445
26.7199 | -100.856
26.7236 | -100.868
26.7273 | -100.8817
26.731 | -100.8936
26.7347 | -100.9014
26.7384 | -100.908
26.7421 | -100.9163
26.7458 | -100.9252
26.7496 | -100.933
26.7533 | -100.94
26.757 | -100.9472
26.7607 | -100.9555
26.7644 | -100.9622
26.7681 | -100.9679
26.7718 | -100.9648
26.7755 | -100.9705
26.7793 | -100.9772
26.783 | -100.9809
26.7867 | -100.9834
26.7904 | -100.9839
26.7941 | -100.9857
26.7978 | -100.9867
26.8015 | -100.988
26.8052 | -100.9881
26.8089 | -100.9875
26.8127 | -100.9858
26.8164 | -100.9822
26.8201 | -100.9807
26.8238 | -100.979
26.8275 | -100.9761
26.8312 | -100.9749
26.8349 | -100.9755
26.8386 | -100.9732
26.8423 | -100.9713
26.8461 | -100.9687
26.8498 | -100.9658
26.8535 | -100.9629
26.8572 | -100.9599
26.8609 | -100.9555
26.8646 | -100.9531
26.8683 | -100.9502
26.872 | -100.9477
26.8758 | -100.9473
26.8795 | -100.9474
26.8832 | -100.947
26.8869 | -100.9457
26.8906 | -100.9455
26.8943 | -100.9444
26.898 | -100.9438
26.9017 | -100.9422
26.9054 | -100.9408
26.9092 | -100.9415
26.9129 | -100.9431
26.9166 | -100.9449
26.9203 | -100.9486
26.924 | -100.9526
26.9277 | -100.9558
26.9314 | -100.961
26.9351 | -100.9651
26.9389 | -100.9685
26.9426 | -100.9734
26.9463 | -100.976
26.95 | -100.9758
26.9537 | -100.9783
26.9574 | -100.983
26.9611 | -100.9868
26.9648 | -100.9906
26.9685 | -100.9958
26.9723 | -101.0019
26.976 | -101.0079
26.9797 | -101.0131
26.9834 | -101.0168
26.9871 | -101.0211
26.9908 | -101.0282
26.9945 | -101.0345
26.9982 | -101.0391
27.0019 | -101.0436
27.0057 | -101.0481
27.0094 | -101.0529
27.0131 | -101.0556
27.0168 | -101.0566
27.0205 | -101.0587
27.0242 | -101.0632
27.0279 | -101.068
27.0316 | -101.0713
27.0354 | -101.0765
27.0391 | -101.0798
27.0428 | -101.0803
27.0465 | -101.0806
27.0502 | -101.0796
27.0539 | -101.0786
27.0576 | -101.0784
27.0613 | -101.0783
27.065 | -101.0767
27.0688 | -101.0757
27.0725 | -101.0759
27.0762 | -101.0737
27.0799 | -101.0704
27.0836 | -101.0674
27.0873 | -101.0646
27.091 | -101.064
27.0947 | -101.064
27.0984 | -101.0623
27.1022 | -101.0597
27.1059 | -101.056
27.1096 | -101.0537
27.1133 | -101.0543
27.117 | -101.0556
27.1207 | -101.0548
27.1244 | -101.053
27.1281 | -101.0503
27.1319 | -101.0475
27.1356 | -101.0457
27.1393 | -101.0443
27.143 | -101.0459
27.1467 | -101.0483
27.1504 | -101.0514
27.1541 | -101.0551
27.1578 | -101.0584
27.1615 | -101.0621
27.1653 | -101.0665
27.169 | -101.0723
27.1727 | -101.0777
27.1764 | -101.0825
27.1801 | -101.0884
27.1838 | -101.0943
27.1875 | -101.1007
27.1912 | -101.108
27.195 | -101.1175
27.1987 | -101.1289
27.2024 | -101.1378
27.2061 | -101.1489
27.2098 | -101.1604
27.2135 | -101.1719
27.2172 | -101.1836
27.2209 | -101.1958
27.2246 | -101.2079
27.2284 | -101.2198
27.2321 | -101.2314
27.2358 | -101.2439
27.2395 | -101.2579
27.2432 | -101.2738
27.2469 | -101.2851
27.2506 | -101.2971
27.2543 | -101.3102
27.2581 | -101.3228
27.2618 | -101.3369
27.2655 | -101.3528
27.2692 | -101.3697
27.2729 | -101.387
27.2766 | -101.4045
27.2803 | -101.4214
27.284 | -101.4402
27.2877 | -101.459
27.2915 | -101.475
27.2952 | -101.4914
27.2989 | -101.5085
27.3026 | -101.5258
27.3063 | -101.5467
27.31 | -101.5637
27.3137 | -101.5817
27.3174 | -101.5993
27.3211 | -101.6186
27.3249 | -101.6407
27.3286 | -101.6618
27.3323 | -101.6852
27.336 | -101.7091
27.3397 | -101.7297
27.3434 | -101.7489
27.3471 | -101.7687
27.3508 | -101.7875
27.3546 | -101.8059
27.3583 | -101.8244
27.362 | -101.8424
27.3657 | -101.8603
27.3694 | -101.8792
27.3731 | -101.8972
27.3768 | -101.9128
27.3805 | -101.9297
27.3842 | -101.9465
27.388 | -101.9615
27.3917 | -101.9766
27.3954 | -101.9923
27.3991 | -102.0093
27.4028 | -102.0251
27.4065 | -102.0372
27.4102 | -102.0494
27.4139 | -102.0604
27.4176 | -102.071
27.4214 | -102.0811
27.4251 | -102.0911
27.4288 | -102.1
27.4325 | -102.1078
27.4362 | -102.1162
27.4399 | -102.1231
27.4436 | -102.1299
27.4473 | -102.1368
27.4511 | -102.1429
27.4548 | -102.1488
27.4585 | -102.1556
27.4622 | -102.1615
27.4659 | -102.1669
27.4696 | -102.1716
27.4733 | -102.1755
27.477 | -102.1781
27.4807 | -102.1782
27.4845 | -102.1782
27.4882 | -102.1791
27.4919 | -102.1796
27.4956 | -102.1794
27.4993 | -102.1804
27.503 | -102.1824
27.5067 | -102.1837
27.5104 | -102.183
27.5142 | -102.1821
27.5179 | -102.1814
27.5216 | -102.1809
27.5253 | -102.1808
27.529 | -102.1787
27.5327 | -102.1773
27.5364 | -102.1772
27.5401 | -102.1776
27.5438 | -102.1773
27.5476 | -102.1768
27.5513 | -102.1764
27.555 | -102.1744
27.5587 | -102.1746
27.5624 | -102.1756
27.5661 | -102.1762
27.5698 | -102.1782
27.5735 | -102.1804
27.5772 | -102.1838
27.581 | -102.1849
27.5847 | -102.1861
27.5884 | -102.1895
27.5921 | -102.1928
27.5958 | -102.1951
27.5995 | -102.1972
27.6032 | -102.1979
27.6069 | -102.2006
27.6107 | -102.2019
27.6144 | -102.2037
27.6181 | -102.2066
27.6218 | -102.2081
27.6255 | -102.211
27.6292 | -102.2127
27.6329 | -102.2153
27.6366 | -102.2182
27.6403 | -102.2222
27.6441 | -102.2243
27.6478 | -102.2236
27.6515 | -102.2218
27.6552 | -102.2199
27.6589 | -102.2177
27.6626 | -102.2172
27.6663 | -102.2163
27.67 | -102.2131
27.6737 | -102.2098
27.6775 | -102.2085
27.6812 | -102.2062
27.6849 | -102.2018
27.6886 | -102.1984
27.6923 | -102.1922
27.696 | -102.1854
27.6997 | -102.1802
27.7034 | -102.1736
27.7072 | -102.1665
27.7109 | -102.1613
27.7146 | -102.1533
27.7183 | -102.144
27.722 | -102.1352
27.7257 | -102.1256
27.7294 | -102.1134
27.7331 | -102.1015
27.7368 | -102.09
27.7406 | -102.0778
27.7443 | -102.0654
27.748 | -102.0527
27.7517 | -102.0413
27.7554 | -102.0298
27.7591 | -102.0178
27.7628 | -102.0058
27.7665 | -101.9951
27.7703 | -101.9836
27.774 | -101.9703
27.7777 | -101.9582
27.7814 | -101.9456
27.7851 | -101.9319
27.7888 | -101.9175
27.7925 | -101.904
27.7962 | -101.8917
27.7999 | -101.8792
27.8037 | -101.8682
27.8074 | -101.8565
27.8111 | -101.8448
27.8148 | -101.8323
27.8185 | -101.8207
27.8222 | -101.8108
27.8259 | -101.802
27.8296 | -101.7937
27.8333 | -101.7839
27.8371 | -101.7756
27.8408 | -101.7729
27.8445 | -101.7645
27.8482 | -101.7575
27.8519 | -101.7498
27.8556 | -101.7387
27.8593 | -101.727
27.863 | -101.7201
27.8668 | -101.7132
27.8705 | -101.7066
27.8742 | -101.7019
27.8779 | -101.6978
27.8816 | -101.6936
27.8853 | -101.6903
27.889 | -101.6864
27.8927 | -101.6803
27.8964 | -101.6756
27.9002 | -101.6723
27.9039 | -101.6695
27.9076 | -101.6675
27.9113 | -101.6659
27.915 | -101.6626
27.9187 | -101.6586
27.9224 | -101.6524
27.9261 | -101.6357
27.9299 | -101.6162
27.9336 | -101.605
27.9373 | -101.5973
27.941 | -101.5902
27.9447 | -101.5815
27.9484 | -101.573
27.9521 | -101.5646
27.9558 | -101.5561
27.9595 | -101.5469
27.9633 | -101.5364
27.967 | -101.5257
27.9707 | -101.5122
27.9744 | -101.4994
27.9781 | -101.4876
27.9818 | -101.473
27.9855 | -101.4575
27.9892 | -101.4433
27.9929 | -101.4269
27.9967 | -101.4068
28.0004 | -101.3866