## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 257 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 243
Hood | 02-92B/P | 188
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1012.6 | hPa |  | PRC PRO 01
Humidity | 55.9 | %/Hr
Internal temperature of calibration bench | 23.65 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1020 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.4 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.01 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.06 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3070 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 27/06/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 05/07/2018
Technical Manager | Anthony HUE | Approval |  | 16/07/2017
CEO | Robert GLEMOT | Approval |  | 16/07/2017
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.1611 | 58.1871
0.012209 | -35.8843 | 39.3074
0.018313 | -34.9597 | 28.8351
0.024417 | -34.5212 | 22.0697
0.030521 | -34.2891 | 17.9198
0.036626 | -34.1387 | 14.8959
0.04273 | -34.0598 | 12.9292
0.048834 | -33.9984 | 11.1284
0.054938 | -33.983 | 9.7256
0.061043 | -33.9657 | 8.7397
0.067147 | -33.9464 | 7.8979
0.073251 | -33.9219 | 7.1698
0.079355 | -33.8913 | 6.6449
0.08546 | -33.8716 | 6.2683
0.091564 | -33.8665 | 5.8639
0.097668 | -33.8589 | 5.3701
0.10377 | -33.8462 | 4.9
0.10988 | -33.8406 | 4.4797
0.11598 | -33.8429 | 4.1346
0.12209 | -33.8527 | 3.8908
0.12819 | -33.8651 | 3.7246
0.13429 | -33.875 | 3.5927
0.1404 | -33.8833 | 3.4703
0.1465 | -33.8926 | 3.3349
0.15261 | -33.9021 | 3.1832
0.15871 | -33.9086 | 3.0016
0.16482 | -33.9085 | 2.7774
0.17092 | -33.8903 | 2.568
0.17702 | -33.8766 | 2.3789
0.18313 | -33.8624 | 2.2271
0.18923 | -33.8516 | 2.1146
0.19534 | -33.8474 | 2.0444
0.20144 | -33.8488 | 1.9987
0.20754 | -33.8567 | 1.9728
0.21365 | -33.8611 | 1.9346
0.21975 | -33.8667 | 1.8862
0.22586 | -33.8735 | 1.8272
0.23196 | -33.8787 | 1.7425
0.23807 | -33.8814 | 1.6427
0.24417 | -33.8778 | 1.5321
0.25027 | -33.8754 | 1.4347
0.25638 | -33.8723 | 1.3339
0.26248 | -33.8683 | 1.2503
0.26859 | -33.8654 | 1.1723
0.27469 | -33.8649 | 1.1045
0.2808 | -33.8697 | 1.0721
0.2869 | -33.8736 | 1.0296
0.293 | -33.8782 | 0.99508
0.29911 | -33.8836 | 0.97088
0.30521 | -33.8887 | 0.94952
0.31132 | -33.8923 | 0.92364
0.31742 | -33.8954 | 0.89465
0.32353 | -33.8991 | 0.85907
0.32963 | -33.9013 | 0.81874
0.33573 | -33.9024 | 0.76776
0.34184 | -33.9036 | 0.71712
0.34794 | -33.9041 | 0.65641
0.35405 | -33.9047 | 0.59991
0.36015 | -33.9049 | 0.54375
0.36626 | -33.9048 | 0.48908
0.37236 | -33.9045 | 0.43872
0.37846 | -33.9039 | 0.39495
0.38457 | -33.9022 | 0.37487
0.39067 | -33.9019 | 0.34469
0.39678 | -33.8999 | 0.31865
0.40288 | -33.8975 | 0.30053
0.40899 | -33.8947 | 0.28121
0.41509 | -33.8921 | 0.26117
0.42119 | -33.888 | 0.23589
0.4273 | -33.8851 | 0.21274
0.4334 | -33.8826 | 0.1895
0.43951 | -33.8799 | 0.16326
0.44561 | -33.8777 | 0.13492
0.45172 | -33.8775 | 0.10323
0.45782 | -33.8767 | 0.075515
0.46392 | -33.8767 | 0.050308
0.47003 | -33.8774 | 0.023013
0.47613 | -33.8788 | 0.0022376
0.48224 | -33.881 | -0.015085
0.48834 | -33.8857 | -0.021382
0.49445 | -33.889 | -0.030886
0.50055 | -33.8932 | -0.042841
0.50665 | -33.8969 | -0.046368
0.51276 | -33.9006 | -0.051048
0.51886 | -33.9042 | -0.052622
0.52497 | -33.9061 | -0.057855
0.53107 | -33.9085 | -0.058849
0.53717 | -33.9108 | -0.062386
0.54328 | -33.9127 | -0.068271
0.54938 | -33.9137 | -0.077147
0.55549 | -33.914 | -0.088818
0.56159 | -33.9127 | -0.11631
0.5677 | -33.9124 | -0.13907
0.5738 | -33.9109 | -0.16123
0.5799 | -33.9102 | -0.1841
0.58601 | -33.9095 | -0.21889
0.59211 | -33.9083 | -0.26439
0.59822 | -33.9074 | -0.29962
0.60432 | -33.9065 | -0.33501
0.61043 | -33.9059 | -0.3766
0.61653 | -33.9057 | -0.41414
0.62263 | -33.9048 | -0.45457
0.62874 | -33.9057 | -0.48914
0.63484 | -33.9059 | -0.52734
0.64095 | -33.907 | -0.56363
0.64705 | -33.9081 | -0.59943
0.65316 | -33.9094 | -0.63004
0.65926 | -33.9109 | -0.65788
0.66536 | -33.9129 | -0.67826
0.67147 | -33.9145 | -0.7041
0.67757 | -33.9162 | -0.72941
0.68368 | -33.918 | -0.74882
0.68978 | -33.9201 | -0.76912
0.69589 | -33.9219 | -0.78638
0.70199 | -33.9234 | -0.79997
0.70809 | -33.925 | -0.81839
0.7142 | -33.9264 | -0.83341
0.7203 | -33.928 | -0.85089
0.72641 | -33.929 | -0.86357
0.73251 | -33.9305 | -0.8802
0.73862 | -33.9307 | -0.89612
0.74472 | -33.9322 | -0.91073
0.75082 | -33.9328 | -0.92329
0.75693 | -33.933 | -0.93843
0.76303 | -33.9326 | -0.95195
0.76914 | -33.9306 | -0.97117
0.77524 | -33.9298 | -0.99072
0.78135 | -33.9287 | -1.0077
0.78745 | -33.9276 | -1.0261
0.79355 | -33.9266 | -1.0444
0.79966 | -33.9252 | -1.0642
0.80576 | -33.9231 | -1.0873
0.81187 | -33.9213 | -1.1069
0.81797 | -33.9192 | -1.128
0.82408 | -33.9176 | -1.1467
0.83018 | -33.9158 | -1.1693
0.83628 | -33.9141 | -1.1944
0.84239 | -33.9125 | -1.2137
0.84849 | -33.9107 | -1.2313
0.8546 | -33.9092 | -1.2514
0.8607 | -33.9077 | -1.2714
0.86681 | -33.9066 | -1.2935
0.87291 | -33.905 | -1.3132
0.87901 | -33.9043 | -1.3309
0.88512 | -33.9034 | -1.349
0.89122 | -33.9026 | -1.3672
0.89733 | -33.9022 | -1.3857
0.90343 | -33.9018 | -1.4028
0.90953 | -33.9022 | -1.4188
0.91564 | -33.9022 | -1.4336
0.92174 | -33.902 | -1.448
0.92785 | -33.9019 | -1.4623
0.93395 | -33.9023 | -1.4761
0.94006 | -33.9025 | -1.4932
0.94616 | -33.9034 | -1.5028
0.95226 | -33.9038 | -1.5178
0.95837 | -33.9043 | -1.5311
0.96447 | -33.9049 | -1.5445
0.97058 | -33.9057 | -1.5567
0.97668 | -33.9065 | -1.5669
0.98279 | -33.9075 | -1.5734
0.98889 | -33.9085 | -1.5848
0.99499 | -33.9092 | -1.5951
1.0011 | -33.9102 | -1.6064
1.0072 | -33.9112 | -1.6148
1.0133 | -33.9122 | -1.6249
1.0194 | -33.9134 | -1.6323
1.0255 | -33.9141 | -1.6402
1.0316 | -33.9149 | -1.6483
1.0377 | -33.9156 | -1.655
1.0438 | -33.9164 | -1.662
1.0499 | -33.9166 | -1.666
1.056 | -33.9172 | -1.6706
1.0621 | -33.9175 | -1.6759
1.0682 | -33.9183 | -1.6834
1.0743 | -33.9188 | -1.6887
1.0805 | -33.9195 | -1.6946
1.0866 | -33.9199 | -1.7029
1.0927 | -33.9204 | -1.7092
1.0988 | -33.9212 | -1.718
1.1049 | -33.9215 | -1.7248
1.111 | -33.9218 | -1.7331
1.1171 | -33.9221 | -1.7408
1.1232 | -33.922 | -1.7518
1.1293 | -33.922 | -1.7628
1.1354 | -33.922 | -1.7726
1.1415 | -33.9219 | -1.7832
1.1476 | -33.9221 | -1.7937
1.1537 | -33.922 | -1.8048
1.1598 | -33.9222 | -1.8203
1.1659 | -33.9225 | -1.8323
1.172 | -33.9226 | -1.8445
1.1781 | -33.9228 | -1.8574
1.1842 | -33.9229 | -1.8715
1.1903 | -33.9235 | -1.8901
1.1964 | -33.924 | -1.9042
1.2025 | -33.9242 | -1.9181
1.2086 | -33.9245 | -1.9348
1.2147 | -33.9252 | -1.9508
1.2209 | -33.9257 | -1.9673
1.227 | -33.9267 | -1.9859
1.2331 | -33.9275 | -2.0019
1.2392 | -33.9282 | -2.0197
1.2453 | -33.9291 | -2.0349
1.2514 | -33.9301 | -2.0522
1.2575 | -33.9311 | -2.0696
1.2636 | -33.9324 | -2.0876
1.2697 | -33.9335 | -2.1055
1.2758 | -33.9344 | -2.1226
1.2819 | -33.9353 | -2.141
1.288 | -33.9363 | -2.1583
1.2941 | -33.9373 | -2.1762
1.3002 | -33.9385 | -2.1924
1.3063 | -33.9396 | -2.2088
1.3124 | -33.9407 | -2.2251
1.3185 | -33.9418 | -2.2421
1.3246 | -33.9428 | -2.259
1.3307 | -33.9439 | -2.2755
1.3368 | -33.9447 | -2.2895
1.3429 | -33.9456 | -2.3048
1.349 | -33.9466 | -2.319
1.3551 | -33.9475 | -2.3355
1.3613 | -33.9485 | -2.3509
1.3674 | -33.9491 | -2.3635
1.3735 | -33.9499 | -2.3765
1.3796 | -33.9508 | -2.39
1.3857 | -33.9516 | -2.4027
1.3918 | -33.9521 | -2.4165
1.3979 | -33.9528 | -2.4307
1.404 | -33.953 | -2.4425
1.4101 | -33.9536 | -2.4552
1.4162 | -33.9541 | -2.4671
1.4223 | -33.9544 | -2.4793
1.4284 | -33.9548 | -2.4918
1.4345 | -33.9554 | -2.5037
1.4406 | -33.9557 | -2.5128
1.4467 | -33.956 | -2.5236
1.4528 | -33.9563 | -2.5346
1.4589 | -33.9567 | -2.546
1.465 | -33.9569 | -2.5563
1.4711 | -33.9573 | -2.567
1.4772 | -33.9574 | -2.5773
1.4833 | -33.9576 | -2.5883
1.4894 | -33.9578 | -2.5998
1.4955 | -33.958 | -2.6119
1.5016 | -33.9581 | -2.6235
1.5078 | -33.9583 | -2.6351
1.5139 | -33.9587 | -2.6464
1.52 | -33.9589 | -2.6588
1.5261 | -33.9591 | -2.6698
1.5322 | -33.9592 | -2.6794
1.5383 | -33.9592 | -2.6909
1.5444 | -33.9593 | -2.7029
1.5505 | -33.9593 | -2.7139
1.5566 | -33.9594 | -2.7251
1.5627 | -33.9596 | -2.7356
1.5688 | -33.9597 | -2.7461
1.5749 | -33.9598 | -2.7573
1.581 | -33.96 | -2.7714
1.5871 | -33.9603 | -2.784
1.5932 | -33.9605 | -2.7954
1.5993 | -33.9606 | -2.8084
1.6054 | -33.9607 | -2.8211
1.6115 | -33.9607 | -2.8343
1.6176 | -33.961 | -2.8492
1.6237 | -33.9612 | -2.8607
1.6298 | -33.9614 | -2.8748
1.6359 | -33.9616 | -2.8877
1.642 | -33.9618 | -2.9019
1.6482 | -33.9622 | -2.9165
1.6543 | -33.9625 | -2.93
1.6604 | -33.9627 | -2.9442
1.6665 | -33.9631 | -2.958
1.6726 | -33.9634 | -2.9729
1.6787 | -33.9637 | -2.9883
1.6848 | -33.964 | -3.0032
1.6909 | -33.9644 | -3.0177
1.697 | -33.9646 | -3.0329
1.7031 | -33.9649 | -3.047
1.7092 | -33.9654 | -3.0611
1.7153 | -33.9656 | -3.0768
1.7214 | -33.966 | -3.0912
1.7275 | -33.9663 | -3.1055
1.7336 | -33.9665 | -3.1196
1.7397 | -33.9667 | -3.1337
1.7458 | -33.9672 | -3.1493
1.7519 | -33.9673 | -3.1645
1.758 | -33.9675 | -3.1786
1.7641 | -33.9677 | -3.1931
1.7702 | -33.9679 | -3.2066
1.7763 | -33.9684 | -3.2205
1.7824 | -33.9686 | -3.2353
1.7885 | -33.9688 | -3.2481
1.7947 | -33.9691 | -3.2622
1.8008 | -33.9693 | -3.2756
1.8069 | -33.9696 | -3.2892
1.813 | -33.9698 | -3.3025
1.8191 | -33.97 | -3.3149
1.8252 | -33.9699 | -3.3269
1.8313 | -33.9701 | -3.3402
1.8374 | -33.9702 | -3.3533
1.8435 | -33.9705 | -3.365
1.8496 | -33.9705 | -3.377
1.8557 | -33.9705 | -3.3898
1.8618 | -33.9704 | -3.4006
1.8679 | -33.9705 | -3.4129
1.874 | -33.9707 | -3.4254
1.8801 | -33.9707 | -3.4371
1.8862 | -33.9707 | -3.4491
1.8923 | -33.9708 | -3.4604
1.8984 | -33.9706 | -3.4705
1.9045 | -33.9707 | -3.4817
1.9106 | -33.9707 | -3.4926
1.9167 | -33.9707 | -3.503
1.9228 | -33.9708 | -3.5136
1.9289 | -33.9706 | -3.5222
1.9351 | -33.9707 | -3.5327
1.9412 | -33.9707 | -3.5435
1.9473 | -33.9706 | -3.5547
1.9534 | -33.9705 | -3.5662
1.9595 | -33.9703 | -3.5775
1.9656 | -33.9698 | -3.5871
1.9717 | -33.9695 | -3.5974
1.9778 | -33.9693 | -3.6093
1.9839 | -33.9693 | -3.621
1.99 | -33.9693 | -3.6319
1.9961 | -33.9693 | -3.6415
2.0022 | -33.9692 | -3.6517
2.0083 | -33.969 | -3.6624
2.0144 | -33.969 | -3.6725
2.0205 | -33.9688 | -3.684
2.0266 | -33.9687 | -3.6958
2.0327 | -33.9687 | -3.7069
2.0388 | -33.9686 | -3.7205
2.0449 | -33.9687 | -3.7313
2.051 | -33.9687 | -3.7416
2.0571 | -33.9683 | -3.7532
2.0632 | -33.9682 | -3.764
2.0693 | -33.9683 | -3.7741
2.0754 | -33.9685 | -3.7861
2.0816 | -33.9684 | -3.797
2.0877 | -33.9684 | -3.8081
2.0938 | -33.9684 | -3.8194
2.0999 | -33.9684 | -3.8307
2.106 | -33.9688 | -3.8423
2.1121 | -33.9688 | -3.8537
2.1182 | -33.9689 | -3.865
2.1243 | -33.969 | -3.8767
2.1304 | -33.9691 | -3.8877
2.1365 | -33.9693 | -3.8987
2.1426 | -33.9695 | -3.911
2.1487 | -33.9699 | -3.9229
2.1548 | -33.97 | -3.9352
2.1609 | -33.9702 | -3.9472
2.167 | -33.9704 | -3.96
2.1731 | -33.9706 | -3.971
2.1792 | -33.9714 | -3.9829
2.1853 | -33.9716 | -3.9948
2.1914 | -33.972 | -4.0071
2.1975 | -33.9723 | -4.019
2.2036 | -33.9728 | -4.0312
2.2097 | -33.9732 | -4.0432
2.2158 | -33.9739 | -4.0547
2.222 | -33.9744 | -4.0664
2.2281 | -33.9747 | -4.0773
2.2342 | -33.9751 | -4.0904
2.2403 | -33.9756 | -4.1018
2.2464 | -33.9763 | -4.1138
2.2525 | -33.9767 | -4.1246
2.2586 | -33.9774 | -4.1358
2.2647 | -33.9781 | -4.1467
2.2708 | -33.9786 | -4.1585
2.2769 | -33.9792 | -4.1711
2.283 | -33.9798 | -4.1808
2.2891 | -33.9804 | -4.1919
2.2952 | -33.9811 | -4.2031
2.3013 | -33.9817 | -4.2141
2.3074 | -33.9824 | -4.2252
2.3135 | -33.9829 | -4.2366
2.3196 | -33.9835 | -4.2471
2.3257 | -33.9839 | -4.259
2.3318 | -33.9844 | -4.2686
2.3379 | -33.9849 | -4.281
2.344 | -33.9854 | -4.2919
2.3501 | -33.9859 | -4.3037
2.3562 | -33.9863 | -4.3137
2.3623 | -33.9867 | -4.324
2.3685 | -33.9872 | -4.3345
2.3746 | -33.9876 | -4.3452
2.3807 | -33.9881 | -4.3556
2.3868 | -33.9883 | -4.3655
2.3929 | -33.9888 | -4.376
2.399 | -33.9893 | -4.3856
2.4051 | -33.9897 | -4.3965
2.4112 | -33.9901 | -4.4066
2.4173 | -33.9904 | -4.4167
2.4234 | -33.9907 | -4.4262
2.4295 | -33.991 | -4.4356
2.4356 | -33.9912 | -4.4458
2.4417 | -33.9914 | -4.4562
2.4478 | -33.9917 | -4.4666
2.4539 | -33.9921 | -4.4763
2.46 | -33.9919 | -4.4858
2.4661 | -33.992 | -4.4954
2.4722 | -33.9923 | -4.5061
2.4783 | -33.9926 | -4.516
2.4844 | -33.9929 | -4.5251
2.4905 | -33.993 | -4.535
2.4966 | -33.9927 | -4.5448
2.5027 | -33.9929 | -4.5542
2.5089 | -33.9929 | -4.5634
2.515 | -33.993 | -4.5737
2.5211 | -33.9931 | -4.5829
2.5272 | -33.993 | -4.5925
2.5333 | -33.9929 | -4.6027
2.5394 | -33.9929 | -4.6125
2.5455 | -33.9929 | -4.6219
2.5516 | -33.9929 | -4.6321
2.5577 | -33.9928 | -4.6418
2.5638 | -33.9927 | -4.6512
2.5699 | -33.9927 | -4.6614
2.576 | -33.9928 | -4.6713
2.5821 | -33.9927 | -4.6813
2.5882 | -33.9926 | -4.6911
2.5943 | -33.9926 | -4.7017
2.6004 | -33.9924 | -4.7121
2.6065 | -33.9924 | -4.722
2.6126 | -33.9924 | -4.7325
2.6187 | -33.9924 | -4.7424
2.6248 | -33.9923 | -4.7523
2.6309 | -33.9922 | -4.7631
2.637 | -33.9922 | -4.7736
2.6431 | -33.9922 | -4.7844
2.6492 | -33.9922 | -4.7942
2.6554 | -33.9922 | -4.8051
2.6615 | -33.9921 | -4.8154
2.6676 | -33.992 | -4.8254
2.6737 | -33.992 | -4.8362
2.6798 | -33.9919 | -4.8461
2.6859 | -33.9919 | -4.8565
2.692 | -33.9918 | -4.8668
2.6981 | -33.9918 | -4.878
2.7042 | -33.9919 | -4.8901
2.7103 | -33.9918 | -4.901
2.7164 | -33.9919 | -4.912
2.7225 | -33.9919 | -4.923
2.7286 | -33.9919 | -4.9344
2.7347 | -33.9919 | -4.9454
2.7408 | -33.992 | -4.9563
2.7469 | -33.992 | -4.9675
2.753 | -33.9921 | -4.979
2.7591 | -33.9921 | -4.9906
2.7652 | -33.992 | -5.0015
2.7713 | -33.9921 | -5.0121
2.7774 | -33.9924 | -5.0237
2.7835 | -33.9926 | -5.0347
2.7896 | -33.9927 | -5.045
2.7958 | -33.993 | -5.0559
2.8019 | -33.9933 | -5.0667
2.808 | -33.9935 | -5.0791
2.8141 | -33.9937 | -5.0917
2.8202 | -33.9939 | -5.1042
2.8263 | -33.9942 | -5.1153
2.8324 | -33.9942 | -5.1272
2.8385 | -33.9943 | -5.1394
2.8446 | -33.9947 | -5.1515
2.8507 | -33.9947 | -5.1639
2.8568 | -33.9949 | -5.176
2.8629 | -33.9953 | -5.1881
2.869 | -33.9955 | -5.1994
2.8751 | -33.9957 | -5.2119
2.8812 | -33.9961 | -5.2245
2.8873 | -33.9963 | -5.236
2.8934 | -33.9964 | -5.2479
2.8995 | -33.9966 | -5.2597
2.9056 | -33.9968 | -5.2703
2.9117 | -33.9971 | -5.2818
2.9178 | -33.9975 | -5.2948
2.9239 | -33.9976 | -5.3071
2.93 | -33.9978 | -5.3198
2.9361 | -33.9982 | -5.3316
2.9423 | -33.9986 | -5.3442
2.9484 | -33.9987 | -5.3574
2.9545 | -33.9992 | -5.3705
2.9606 | -33.9994 | -5.3831
2.9667 | -33.9996 | -5.3948
2.9728 | -33.9999 | -5.4064
2.9789 | -34 | -5.4188
2.985 | -34.0001 | -5.4319
2.9911 | -34.0003 | -5.4439
2.9972 | -34.0005 | -5.4562
3.0033 | -34.0008 | -5.4683
3.0094 | -34.001 | -5.4811
3.0155 | -34.0013 | -5.4936
3.0216 | -34.0015 | -5.5064
3.0277 | -34.0018 | -5.519
3.0338 | -34.0021 | -5.5317
3.0399 | -34.0024 | -5.5446
3.046 | -34.0026 | -5.5566
3.0521 | -34.0029 | -5.5689
3.0582 | -34.0031 | -5.5809
3.0643 | -34.0035 | -5.5931
3.0704 | -34.0036 | -5.605
3.0765 | -34.0039 | -5.6171
3.0827 | -34.0041 | -5.6297
3.0888 | -34.0042 | -5.642
3.0949 | -34.0043 | -5.6542
3.101 | -34.0045 | -5.6673
3.1071 | -34.0047 | -5.6801
3.1132 | -34.0049 | -5.6926
3.1193 | -34.0051 | -5.7047
3.1254 | -34.0052 | -5.716
3.1315 | -34.0054 | -5.7285
3.1376 | -34.0056 | -5.7406
3.1437 | -34.0057 | -5.753
3.1498 | -34.0059 | -5.7652
3.1559 | -34.0061 | -5.7774
3.162 | -34.0061 | -5.7896
3.1681 | -34.0061 | -5.8021
3.1742 | -34.0063 | -5.8147
3.1803 | -34.0066 | -5.8266
3.1864 | -34.0067 | -5.8382
3.1925 | -34.0069 | -5.8508
3.1986 | -34.0069 | -5.8619
3.2047 | -34.007 | -5.874
3.2108 | -34.0072 | -5.8855
3.2169 | -34.0075 | -5.8979
3.223 | -34.0074 | -5.9101
3.2292 | -34.0074 | -5.9226
3.2353 | -34.0074 | -5.9323
3.2414 | -34.0075 | -5.9444
3.2475 | -34.0075 | -5.9547
3.2536 | -34.0076 | -5.9675
3.2597 | -34.0078 | -5.9787
3.2658 | -34.008 | -5.9894
3.2719 | -34.0081 | -6.0003
3.278 | -34.0081 | -6.0113
3.2841 | -34.0082 | -6.0229
3.2902 | -34.0083 | -6.0345
3.2963 | -34.0084 | -6.0452
3.3024 | -34.0085 | -6.0552
3.3085 | -34.0086 | -6.0666
3.3146 | -34.0087 | -6.0772
3.3207 | -34.0088 | -6.0883
3.3268 | -34.009 | -6.0996
3.3329 | -34.0091 | -6.1098
3.339 | -34.0093 | -6.1197
3.3451 | -34.0094 | -6.1307
3.3512 | -34.0096 | -6.1406
3.3573 | -34.0097 | -6.1513
3.3634 | -34.0099 | -6.1625
3.3696 | -34.0099 | -6.1736
3.3757 | -34.0101 | -6.1838
3.3818 | -34.0103 | -6.1955
3.3879 | -34.0103 | -6.2056
3.394 | -34.0105 | -6.2154
3.4001 | -34.0105 | -6.2264
3.4062 | -34.0108 | -6.2367
3.4123 | -34.0109 | -6.2471
3.4184 | -34.011 | -6.2574
3.4245 | -34.0111 | -6.2681
3.4306 | -34.0112 | -6.2783
3.4367 | -34.0114 | -6.2886
3.4428 | -34.0118 | -6.2986
3.4489 | -34.0119 | -6.3097
3.455 | -34.0119 | -6.3202
3.4611 | -34.012 | -6.3308
3.4672 | -34.0121 | -6.3418
3.4733 | -34.0123 | -6.3516
3.4794 | -34.0124 | -6.3622
3.4855 | -34.0127 | -6.3727
3.4916 | -34.0129 | -6.3832
3.4977 | -34.013 | -6.3935
3.5038 | -34.0132 | -6.4036
3.5099 | -34.0135 | -6.4143
3.5161 | -34.0137 | -6.4245
3.5222 | -34.014 | -6.4354
3.5283 | -34.0141 | -6.4457
3.5344 | -34.0143 | -6.4565
3.5405 | -34.0145 | -6.467
3.5466 | -34.0147 | -6.4766
3.5527 | -34.015 | -6.4873
3.5588 | -34.0153 | -6.4977
3.5649 | -34.0156 | -6.5082
3.571 | -34.0159 | -6.5188
3.5771 | -34.0161 | -6.5292
3.5832 | -34.0164 | -6.5404
3.5893 | -34.0166 | -6.5506
3.5954 | -34.0168 | -6.5612
3.6015 | -34.017 | -6.5715
3.6076 | -34.0173 | -6.5816
3.6137 | -34.0175 | -6.5929
3.6198 | -34.0178 | -6.6027
3.6259 | -34.018 | -6.6133
3.632 | -34.0183 | -6.6242
3.6381 | -34.0186 | -6.6342
3.6442 | -34.0189 | -6.6451
3.6503 | -34.0192 | -6.6552
3.6565 | -34.0195 | -6.6669
3.6626 | -34.0197 | -6.6782
3.6687 | -34.02 | -6.6889
3.6748 | -34.0203 | -6.6995
3.6809 | -34.0206 | -6.7102
3.687 | -34.0209 | -6.7211
3.6931 | -34.0212 | -6.732
3.6992 | -34.0215 | -6.7428
3.7053 | -34.0217 | -6.7534
3.7114 | -34.0221 | -6.7649
3.7175 | -34.0224 | -6.7748
3.7236 | -34.0226 | -6.7849
3.7297 | -34.0227 | -6.7956
3.7358 | -34.023 | -6.8071
3.7419 | -34.0233 | -6.8186
3.748 | -34.0234 | -6.8298
3.7541 | -34.0236 | -6.841
3.7602 | -34.0239 | -6.8521
3.7663 | -34.0242 | -6.8628
3.7724 | -34.0244 | -6.874
3.7785 | -34.0247 | -6.8849
3.7846 | -34.025 | -6.8961
3.7907 | -34.0253 | -6.9067
3.7969 | -34.0255 | -6.9182
3.803 | -34.0258 | -6.9292
3.8091 | -34.0261 | -6.9399
3.8152 | -34.0263 | -6.9513
3.8213 | -34.0265 | -6.9627
3.8274 | -34.0268 | -6.9739
3.8335 | -34.0271 | -6.9852
3.8396 | -34.0274 | -6.9968
3.8457 | -34.0277 | -7.007
3.8518 | -34.028 | -7.0182
3.8579 | -34.0282 | -7.0289
3.864 | -34.0285 | -7.0396
3.8701 | -34.0288 | -7.0501
3.8762 | -34.0291 | -7.0617
3.8823 | -34.0293 | -7.0722
3.8884 | -34.0296 | -7.0827
3.8945 | -34.0299 | -7.0937
3.9006 | -34.0302 | -7.1048
3.9067 | -34.0305 | -7.116
3.9128 | -34.0308 | -7.1272
3.9189 | -34.0311 | -7.1382
3.925 | -34.0314 | -7.149
3.9311 | -34.0317 | -7.1598
3.9372 | -34.032 | -7.171
3.9434 | -34.0323 | -7.1815
3.9495 | -34.0325 | -7.1928
3.9556 | -34.0327 | -7.2044
3.9617 | -34.033 | -7.216
3.9678 | -34.0332 | -7.2271
3.9739 | -34.0336 | -7.2384
3.98 | -34.0339 | -7.2502
3.9861 | -34.0341 | -7.2607
3.9922 | -34.0344 | -7.2725
3.9983 | -34.0346 | -7.2833
4.0044 | -34.0349 | -7.2939
4.0105 | -34.0352 | -7.3048
4.0166 | -34.0355 | -7.3157
4.0227 | -34.0358 | -7.3262
4.0288 | -34.036 | -7.337
4.0349 | -34.0364 | -7.3474
4.041 | -34.0366 | -7.3585
4.0471 | -34.0369 | -7.3689
4.0532 | -34.0373 | -7.3806
4.0593 | -34.0375 | -7.392
4.0654 | -34.0377 | -7.4032
4.0715 | -34.038 | -7.4139
4.0776 | -34.0382 | -7.4244
4.0838 | -34.0385 | -7.435
4.0899 | -34.0388 | -7.4469
4.096 | -34.0391 | -7.4574
4.1021 | -34.0394 | -7.4687
4.1082 | -34.0397 | -7.4795
4.1143 | -34.04 | -7.4902
4.1204 | -34.0403 | -7.5016
4.1265 | -34.0406 | -7.5125
4.1326 | -34.041 | -7.5234
4.1387 | -34.0412 | -7.535
4.1448 | -34.0416 | -7.546
4.1509 | -34.0419 | -7.557
4.157 | -34.0421 | -7.5677
4.1631 | -34.0424 | -7.5788
4.1692 | -34.0427 | -7.5899
4.1753 | -34.043 | -7.601
4.1814 | -34.0434 | -7.6116
4.1875 | -34.0436 | -7.6228
4.1936 | -34.0439 | -7.6342
4.1997 | -34.0442 | -7.6453
4.2058 | -34.0444 | -7.6566
4.2119 | -34.0447 | -7.6669
4.218 | -34.045 | -7.6777
4.2241 | -34.0454 | -7.688
4.2303 | -34.0457 | -7.6997
4.2364 | -34.0461 | -7.7114
4.2425 | -34.0464 | -7.7223
4.2486 | -34.0467 | -7.7336
4.2547 | -34.047 | -7.7438
4.2608 | -34.0473 | -7.7539
4.2669 | -34.0477 | -7.7653
4.273 | -34.048 | -7.7762
4.2791 | -34.0482 | -7.7874
4.2852 | -34.0484 | -7.7977
4.2913 | -34.0487 | -7.8079
4.2974 | -34.049 | -7.8187
4.3035 | -34.0492 | -7.8294
4.3096 | -34.0495 | -7.84
4.3157 | -34.0497 | -7.8506
4.3218 | -34.05 | -7.8611
4.3279 | -34.0503 | -7.8718
4.334 | -34.0506 | -7.8824
4.3401 | -34.0509 | -7.8935
4.3462 | -34.0511 | -7.9044
4.3523 | -34.0514 | -7.9152
4.3584 | -34.0517 | -7.9248
4.3645 | -34.0519 | -7.9351
4.3707 | -34.0522 | -7.9455
4.3768 | -34.0525 | -7.9564
4.3829 | -34.0526 | -7.9666
4.389 | -34.053 | -7.9769
4.3951 | -34.0533 | -7.9873
4.4012 | -34.0536 | -7.9971
4.4073 | -34.0539 | -8.0081
4.4134 | -34.0542 | -8.0188
4.4195 | -34.0544 | -8.029
4.4256 | -34.0546 | -8.0394
4.4317 | -34.0549 | -8.0495
4.4378 | -34.0552 | -8.0603
4.4439 | -34.0554 | -8.0709
4.45 | -34.0557 | -8.0817
4.4561 | -34.0559 | -8.0923
4.4622 | -34.0562 | -8.1025
4.4683 | -34.0564 | -8.1125
4.4744 | -34.0567 | -8.1227
4.4805 | -34.057 | -8.1331
4.4866 | -34.0573 | -8.1429
4.4927 | -34.0575 | -8.1535
4.4988 | -34.0577 | -8.1635
4.5049 | -34.058 | -8.1739
4.511 | -34.0582 | -8.1841
4.5172 | -34.0586 | -8.1945
4.5233 | -34.0588 | -8.2052
4.5294 | -34.0591 | -8.2154
4.5355 | -34.0593 | -8.2254
4.5416 | -34.0596 | -8.2352
4.5477 | -34.0598 | -8.2452
4.5538 | -34.0601 | -8.2553
4.5599 | -34.0604 | -8.2655
4.566 | -34.0606 | -8.2758
4.5721 | -34.0609 | -8.2857
4.5782 | -34.0612 | -8.2957
4.5843 | -34.0614 | -8.3058
4.5904 | -34.0617 | -8.3156
4.5965 | -34.062 | -8.3252
4.6026 | -34.0623 | -8.3358
4.6087 | -34.0626 | -8.3456
4.6148 | -34.0629 | -8.3564
4.6209 | -34.0633 | -8.3656
4.627 | -34.0636 | -8.3758
4.6331 | -34.0638 | -8.3857
4.6392 | -34.0641 | -8.3954
4.6453 | -34.0643 | -8.4054
4.6514 | -34.0646 | -8.4154
4.6576 | -34.0648 | -8.4259
4.6637 | -34.0651 | -8.4348
4.6698 | -34.0654 | -8.4443
4.6759 | -34.0656 | -8.4539
4.682 | -34.0658 | -8.4639
4.6881 | -34.0662 | -8.4734
4.6942 | -34.0664 | -8.4834
4.7003 | -34.0667 | -8.4933
4.7064 | -34.0669 | -8.5034
4.7125 | -34.0673 | -8.5128
4.7186 | -34.0675 | -8.5223
4.7247 | -34.0678 | -8.5322
4.7308 | -34.0681 | -8.5421
4.7369 | -34.0684 | -8.552
4.743 | -34.0686 | -8.5622
4.7491 | -34.0688 | -8.5718
4.7552 | -34.069 | -8.5814
4.7613 | -34.0693 | -8.5913
4.7674 | -34.0697 | -8.6007
4.7735 | -34.0699 | -8.6106
4.7796 | -34.07 | -8.6202
4.7857 | -34.0702 | -8.6301
4.7918 | -34.0704 | -8.6396
4.7979 | -34.0707 | -8.6492
4.8041 | -34.0709 | -8.6588
4.8102 | -34.0712 | -8.6688
4.8163 | -34.0715 | -8.6788
4.8224 | -34.0716 | -8.6879
4.8285 | -34.0719 | -8.6979
4.8346 | -34.0721 | -8.7081
4.8407 | -34.0725 | -8.7177
4.8468 | -34.0726 | -8.7273
4.8529 | -34.0729 | -8.7371
4.859 | -34.0732 | -8.7464
4.8651 | -34.0734 | -8.7565
4.8712 | -34.0737 | -8.7665
4.8773 | -34.074 | -8.7765
4.8834 | -34.0742 | -8.7866
4.8895 | -34.0745 | -8.7973
4.8956 | -34.0748 | -8.8077
4.9017 | -34.075 | -8.8179
4.9078 | -34.0753 | -8.8276
4.9139 | -34.0755 | -8.8371
4.92 | -34.0757 | -8.8469
4.9261 | -34.076 | -8.8567
4.9322 | -34.0763 | -8.8662
4.9383 | -34.0766 | -8.8757
4.9445 | -34.0769 | -8.8862
4.9506 | -34.0772 | -8.8957
4.9567 | -34.0775 | -8.9056
4.9628 | -34.0778 | -8.9154
4.9689 | -34.078 | -8.9249
4.975 | -34.0782 | -8.9342
4.9811 | -34.0785 | -8.9449
4.9872 | -34.0787 | -8.9552
4.9933 | -34.079 | -8.9646
4.9994 | -34.0793 | -8.9742
5.0055 | -34.0796 | -8.9841
5.0116 | -34.08 | -8.994
5.0177 | -34.0803 | -9.0038
5.0238 | -34.0805 | -9.014
5.0299 | -34.0809 | -9.0246
5.036 | -34.0811 | -9.0345
5.0421 | -34.0813 | -9.0448
5.0482 | -34.0816 | -9.0553
5.0543 | -34.0819 | -9.0652
5.0604 | -34.0822 | -9.0757
5.0665 | -34.0825 | -9.086
5.0726 | -34.0828 | -9.0961
5.0787 | -34.0831 | -9.1064
5.0848 | -34.0833 | -9.1163
5.091 | -34.0836 | -9.1268
5.0971 | -34.0839 | -9.1377
5.1032 | -34.0842 | -9.1477
5.1093 | -34.0845 | -9.1578
5.1154 | -34.0848 | -9.1684
5.1215 | -34.0851 | -9.179
5.1276 | -34.0854 | -9.1893
5.1337 | -34.0857 | -9.1994
5.1398 | -34.0859 | -9.2099
5.1459 | -34.0862 | -9.22
5.152 | -34.0865 | -9.2309
5.1581 | -34.0868 | -9.2415
5.1642 | -34.087 | -9.2516
5.1703 | -34.0874 | -9.2619
5.1764 | -34.0876 | -9.2722
5.1825 | -34.0879 | -9.2831
5.1886 | -34.0881 | -9.2936
5.1947 | -34.0884 | -9.3046
5.2008 | -34.0887 | -9.3145
5.2069 | -34.0891 | -9.3256
5.213 | -34.0894 | -9.3357
5.2191 | -34.0897 | -9.346
5.2252 | -34.09 | -9.3564
5.2314 | -34.0902 | -9.367
5.2375 | -34.0906 | -9.3775
5.2436 | -34.0909 | -9.3878
5.2497 | -34.0912 | -9.3984
5.2558 | -34.0915 | -9.4089
5.2619 | -34.0918 | -9.4195
5.268 | -34.0921 | -9.4298
5.2741 | -34.0924 | -9.4409
5.2802 | -34.0927 | -9.4513
5.2863 | -34.093 | -9.4617
5.2924 | -34.0932 | -9.4727
5.2985 | -34.0936 | -9.484
5.3046 | -34.094 | -9.4952
5.3107 | -34.0943 | -9.5057
5.3168 | -34.0947 | -9.5162
5.3229 | -34.095 | -9.5262
5.329 | -34.0954 | -9.5362
5.3351 | -34.0958 | -9.547
5.3412 | -34.0961 | -9.5574
5.3473 | -34.0965 | -9.5683
5.3534 | -34.0968 | -9.5789
5.3595 | -34.0971 | -9.5898
5.3656 | -34.0974 | -9.6008
5.3717 | -34.0977 | -9.6109
5.3779 | -34.0981 | -9.6212
5.384 | -34.0983 | -9.632
5.3901 | -34.0986 | -9.6423
5.3962 | -34.0989 | -9.6534
5.4023 | -34.0992 | -9.6637
5.4084 | -34.0995 | -9.6742
5.4145 | -34.0998 | -9.6856
5.4206 | -34.1 | -9.6969
5.4267 | -34.1004 | -9.7077
5.4328 | -34.1007 | -9.7178
5.4389 | -34.101 | -9.7286
5.445 | -34.1013 | -9.7394
5.4511 | -34.1017 | -9.7499
5.4572 | -34.102 | -9.7609
5.4633 | -34.1023 | -9.7713
5.4694 | -34.1026 | -9.7818
5.4755 | -34.103 | -9.7931
5.4816 | -34.1032 | -9.8036
5.4877 | -34.1035 | -9.8143
5.4938 | -34.1038 | -9.8252
5.4999 | -34.1041 | -9.8361
5.506 | -34.1044 | -9.8467
5.5121 | -34.1047 | -9.858
5.5183 | -34.105 | -9.8686
5.5244 | -34.1053 | -9.8791
5.5305 | -34.1056 | -9.8901
5.5366 | -34.1059 | -9.9009
5.5427 | -34.1062 | -9.912
5.5488 | -34.1065 | -9.9229
5.5549 | -34.1068 | -9.9337
5.561 | -34.1071 | -9.9445
5.5671 | -34.1075 | -9.9552
5.5732 | -34.1077 | -9.9658
5.5793 | -34.1081 | -9.9769
5.5854 | -34.1084 | -9.9878
5.5915 | -34.1086 | -9.9984
5.5976 | -34.109 | -10.0093
5.6037 | -34.1093 | -10.0198
5.6098 | -34.1096 | -10.0305
5.6159 | -34.1099 | -10.0416
5.622 | -34.1102 | -10.052
5.6281 | -34.1105 | -10.0626
5.6342 | -34.1109 | -10.0729
5.6403 | -34.1111 | -10.0831
5.6464 | -34.1115 | -10.0936
5.6525 | -34.1118 | -10.1039
5.6586 | -34.1122 | -10.1142
5.6648 | -34.1124 | -10.1253
5.6709 | -34.1127 | -10.136
5.677 | -34.113 | -10.1473
5.6831 | -34.1133 | -10.1575
5.6892 | -34.1136 | -10.1679
5.6953 | -34.1139 | -10.1786
5.7014 | -34.1142 | -10.1897
5.7075 | -34.1146 | -10.2007
5.7136 | -34.1149 | -10.2113
5.7197 | -34.1152 | -10.2219
5.7258 | -34.1155 | -10.2327
5.7319 | -34.1158 | -10.2432
5.738 | -34.1161 | -10.2538
5.7441 | -34.1164 | -10.2645
5.7502 | -34.1167 | -10.2751
5.7563 | -34.117 | -10.2855
5.7624 | -34.1173 | -10.2965
5.7685 | -34.1176 | -10.3068
5.7746 | -34.1178 | -10.3169
5.7807 | -34.1182 | -10.3272
5.7868 | -34.1185 | -10.3375
5.7929 | -34.1188 | -10.3482
5.799 | -34.1192 | -10.359
5.8052 | -34.1194 | -10.3699
5.8113 | -34.1197 | -10.3807
5.8174 | -34.12 | -10.3918
5.8235 | -34.1202 | -10.4023
5.8296 | -34.1206 | -10.4133
5.8357 | -34.1208 | -10.4242
5.8418 | -34.1212 | -10.4348
5.8479 | -34.1214 | -10.4456
5.854 | -34.1217 | -10.4563
5.8601 | -34.122 | -10.4669
5.8662 | -34.1223 | -10.4775
5.8723 | -34.1226 | -10.4878
5.8784 | -34.1229 | -10.4986
5.8845 | -34.1232 | -10.5095
5.8906 | -34.1235 | -10.5198
5.8967 | -34.1238 | -10.5304
5.9028 | -34.1241 | -10.5412
5.9089 | -34.1243 | -10.5513
5.915 | -34.1246 | -10.5622
5.9211 | -34.125 | -10.573
5.9272 | -34.1252 | -10.5839
5.9333 | -34.1255 | -10.594
5.9394 | -34.1258 | -10.6045
5.9455 | -34.1261 | -10.6149
5.9517 | -34.1264 | -10.6256
5.9578 | -34.1266 | -10.6362
5.9639 | -34.1269 | -10.6471
5.97 | -34.1272 | -10.6575
5.9761 | -34.1275 | -10.6679
5.9822 | -34.1278 | -10.6784
5.9883 | -34.1281 | -10.6887
5.9944 | -34.1284 | -10.699
6.0005 | -34.1288 | -10.7096
6.0066 | -34.129 | -10.7201
6.0127 | -34.1293 | -10.7303
6.0188 | -34.1296 | -10.7408
6.0249 | -34.1299 | -10.7512
6.031 | -34.1302 | -10.7619
6.0371 | -34.1306 | -10.7726
6.0432 | -34.1309 | -10.7833
6.0493 | -34.1313 | -10.7937
6.0554 | -34.1315 | -10.8044
6.0615 | -34.1318 | -10.8153
6.0676 | -34.1321 | -10.8261
6.0737 | -34.1324 | -10.8365
6.0798 | -34.1327 | -10.847
6.0859 | -34.133 | -10.8574
6.0921 | -34.1334 | -10.8682
6.0982 | -34.1337 | -10.879
6.1043 | -34.134 | -10.8894
6.1104 | -34.1344 | -10.9001
6.1165 | -34.1346 | -10.9105
6.1226 | -34.135 | -10.9206
6.1287 | -34.1353 | -10.9313
6.1348 | -34.1357 | -10.9418
6.1409 | -34.136 | -10.9523
6.147 | -34.1363 | -10.9628
6.1531 | -34.1367 | -10.9732
6.1592 | -34.137 | -10.9839
6.1653 | -34.1373 | -10.9945
6.1714 | -34.1376 | -11.0049
6.1775 | -34.138 | -11.0154
6.1836 | -34.1382 | -11.0258
6.1897 | -34.1386 | -11.0362
6.1958 | -34.1389 | -11.0466
6.2019 | -34.1392 | -11.0571
6.208 | -34.1395 | -11.068
6.2141 | -34.1398 | -11.0787
6.2202 | -34.1401 | -11.0894
6.2263 | -34.1405 | -11.0998
6.2325 | -34.1409 | -11.1104
6.2386 | -34.1412 | -11.121
6.2447 | -34.1415 | -11.1313
6.2508 | -34.1418 | -11.1418
6.2569 | -34.1421 | -11.1525
6.263 | -34.1425 | -11.1628
6.2691 | -34.1429 | -11.1734
6.2752 | -34.1433 | -11.1838
6.2813 | -34.1436 | -11.1944
6.2874 | -34.1439 | -11.2051
6.2935 | -34.1443 | -11.2153
6.2996 | -34.1447 | -11.2261
6.3057 | -34.145 | -11.2364
6.3118 | -34.1454 | -11.2467
6.3179 | -34.1458 | -11.2572
6.324 | -34.1461 | -11.2679
6.3301 | -34.1465 | -11.278
6.3362 | -34.1468 | -11.2886
6.3423 | -34.1472 | -11.2988
6.3484 | -34.1476 | -11.3093
6.3545 | -34.148 | -11.3197
6.3606 | -34.1484 | -11.3303
6.3667 | -34.1488 | -11.3408
6.3728 | -34.1492 | -11.3513
6.379 | -34.1495 | -11.3617
6.3851 | -34.1498 | -11.3726
6.3912 | -34.1503 | -11.383
6.3973 | -34.1507 | -11.3932
6.4034 | -34.1511 | -11.4039
6.4095 | -34.1514 | -11.4146
6.4156 | -34.1517 | -11.4247
6.4217 | -34.1521 | -11.4352
6.4278 | -34.1524 | -11.4456
6.4339 | -34.1529 | -11.456
6.44 | -34.1532 | -11.4663
6.4461 | -34.1536 | -11.477
6.4522 | -34.154 | -11.4872
6.4583 | -34.1543 | -11.4976
6.4644 | -34.1546 | -11.5083
6.4705 | -34.1551 | -11.5193
6.4766 | -34.1555 | -11.5301
6.4827 | -34.1558 | -11.5409
6.4888 | -34.1562 | -11.5516
6.4949 | -34.1566 | -11.562
6.501 | -34.1569 | -11.5726
6.5071 | -34.1573 | -11.5834
6.5132 | -34.1577 | -11.5937
6.5194 | -34.1581 | -11.6039
6.5255 | -34.1584 | -11.6143
6.5316 | -34.1588 | -11.6249
6.5377 | -34.1592 | -11.6356
6.5438 | -34.1596 | -11.6462
6.5499 | -34.16 | -11.6566
6.556 | -34.1604 | -11.667
6.5621 | -34.1608 | -11.6773
6.5682 | -34.1611 | -11.6875
6.5743 | -34.1614 | -11.6986
6.5804 | -34.1618 | -11.7089
6.5865 | -34.1621 | -11.7197
6.5926 | -34.1625 | -11.7302
6.5987 | -34.1628 | -11.7408
6.6048 | -34.1632 | -11.7517
6.6109 | -34.1636 | -11.762
6.617 | -34.164 | -11.7732
6.6231 | -34.1643 | -11.7839
6.6292 | -34.1648 | -11.7942
6.6353 | -34.1651 | -11.8052
6.6414 | -34.1655 | -11.8162
6.6475 | -34.1659 | -11.8267
6.6536 | -34.1664 | -11.837
6.6597 | -34.1668 | -11.8476
6.6659 | -34.1672 | -11.8579
6.672 | -34.1675 | -11.8685
6.6781 | -34.1679 | -11.8788
6.6842 | -34.1683 | -11.8896
6.6903 | -34.1686 | -11.9
6.6964 | -34.169 | -11.9106
6.7025 | -34.1694 | -11.9214
6.7086 | -34.1698 | -11.9317
6.7147 | -34.1702 | -11.9421
6.7208 | -34.1706 | -11.953
6.7269 | -34.1709 | -11.9632
6.733 | -34.1713 | -11.9734
6.7391 | -34.1717 | -11.984
6.7452 | -34.1721 | -11.9943
6.7513 | -34.1725 | -12.0047
6.7574 | -34.1729 | -12.0153
6.7635 | -34.1733 | -12.0257
6.7696 | -34.1737 | -12.0363
6.7757 | -34.1741 | -12.047
6.7818 | -34.1745 | -12.0575
6.7879 | -34.1749 | -12.0682
6.794 | -34.1753 | -12.0788
6.8001 | -34.1757 | -12.0892
6.8063 | -34.1761 | -12.0999
6.8124 | -34.1764 | -12.1102
6.8185 | -34.1768 | -12.1204
6.8246 | -34.1773 | -12.131
6.8307 | -34.1777 | -12.1416
6.8368 | -34.1781 | -12.152
6.8429 | -34.1785 | -12.1623
6.849 | -34.1788 | -12.1727
6.8551 | -34.1792 | -12.1832
6.8612 | -34.1796 | -12.1935
6.8673 | -34.18 | -12.2038
6.8734 | -34.1804 | -12.2146
6.8795 | -34.1808 | -12.2254
6.8856 | -34.1812 | -12.2353
6.8917 | -34.1816 | -12.2451
6.8978 | -34.182 | -12.2559
6.9039 | -34.1823 | -12.2667
6.91 | -34.1827 | -12.2771
6.9161 | -34.183 | -12.2876
6.9222 | -34.1834 | -12.2982
6.9283 | -34.1837 | -12.3085
6.9344 | -34.1841 | -12.3188
6.9405 | -34.1844 | -12.3291
6.9466 | -34.1848 | -12.3393
6.9528 | -34.1851 | -12.3495
6.9589 | -34.1855 | -12.3595
6.965 | -34.1859 | -12.3697
6.9711 | -34.1862 | -12.3798
6.9772 | -34.1866 | -12.3899
6.9833 | -34.187 | -12.4006
6.9894 | -34.1873 | -12.411
6.9955 | -34.1877 | -12.4216
7.0016 | -34.188 | -12.4323
7.0077 | -34.1884 | -12.443
7.0138 | -34.1888 | -12.4537
7.0199 | -34.1891 | -12.4642
7.026 | -34.1894 | -12.4745
7.0321 | -34.1898 | -12.4852
7.0382 | -34.1901 | -12.4957
7.0443 | -34.1905 | -12.5061
7.0504 | -34.1908 | -12.5167
7.0565 | -34.1912 | -12.5273
7.0626 | -34.1915 | -12.538
7.0687 | -34.1918 | -12.5485
7.0748 | -34.1922 | -12.5591
7.0809 | -34.1925 | -12.5697
7.087 | -34.1929 | -12.5799
7.0932 | -34.1933 | -12.5902
7.0993 | -34.1936 | -12.6006
7.1054 | -34.1939 | -12.6108
7.1115 | -34.1942 | -12.6214
7.1176 | -34.1945 | -12.6317
7.1237 | -34.1948 | -12.6423
7.1298 | -34.1951 | -12.6523
7.1359 | -34.1955 | -12.6624
7.142 | -34.1958 | -12.6732
7.1481 | -34.1961 | -12.6836
7.1542 | -34.1965 | -12.6938
7.1603 | -34.1968 | -12.704
7.1664 | -34.1972 | -12.7145
7.1725 | -34.1975 | -12.7245
7.1786 | -34.1978 | -12.7349
7.1847 | -34.1982 | -12.745
7.1908 | -34.1986 | -12.7553
7.1969 | -34.1989 | -12.7654
7.203 | -34.1992 | -12.7761
7.2091 | -34.1995 | -12.7863
7.2152 | -34.1999 | -12.7966
7.2213 | -34.2002 | -12.8071
7.2274 | -34.2005 | -12.8173
7.2335 | -34.2008 | -12.8278
7.2397 | -34.2011 | -12.8382
7.2458 | -34.2014 | -12.8483
7.2519 | -34.2017 | -12.8586
7.258 | -34.2021 | -12.8688
7.2641 | -34.2024 | -12.879
7.2702 | -34.2028 | -12.8894
7.2763 | -34.2031 | -12.8993
7.2824 | -34.2035 | -12.9092
7.2885 | -34.2039 | -12.9194
7.2946 | -34.2042 | -12.9299
7.3007 | -34.2045 | -12.9401
7.3068 | -34.2049 | -12.9504
7.3129 | -34.2052 | -12.9602
7.319 | -34.2055 | -12.9705
7.3251 | -34.2058 | -12.9808
7.3312 | -34.2062 | -12.9907
7.3373 | -34.2065 | -13.0009
7.3434 | -34.2069 | -13.011
7.3495 | -34.2072 | -13.0209
7.3556 | -34.2075 | -13.0314
7.3617 | -34.2078 | -13.0417
7.3678 | -34.2081 | -13.0522
7.3739 | -34.2085 | -13.0619
7.3801 | -34.2088 | -13.0716
7.3862 | -34.2091 | -13.0817
7.3923 | -34.2095 | -13.0916
7.3984 | -34.2098 | -13.1021
7.4045 | -34.2101 | -13.1121
7.4106 | -34.2104 | -13.1221
7.4167 | -34.2107 | -13.1324
7.4228 | -34.211 | -13.142
7.4289 | -34.2114 | -13.152
7.435 | -34.2117 | -13.1623
7.4411 | -34.2121 | -13.1721
7.4472 | -34.2124 | -13.1823
7.4533 | -34.2128 | -13.1927
7.4594 | -34.213 | -13.2028
7.4655 | -34.2134 | -13.2131
7.4716 | -34.2137 | -13.2231
7.4777 | -34.2141 | -13.2328
7.4838 | -34.2144 | -13.2434
7.4899 | -34.2148 | -13.2537
7.496 | -34.2152 | -13.2642
7.5021 | -34.2155 | -13.2743
7.5082 | -34.2158 | -13.2844
7.5143 | -34.2161 | -13.2944
7.5204 | -34.2165 | -13.304
7.5266 | -34.2168 | -13.3144
7.5327 | -34.2171 | -13.3245
7.5388 | -34.2175 | -13.3341
7.5449 | -34.2177 | -13.344
7.551 | -34.218 | -13.3545
7.5571 | -34.2183 | -13.3644
7.5632 | -34.2187 | -13.3742
7.5693 | -34.219 | -13.3846
7.5754 | -34.2193 | -13.3949
7.5815 | -34.2197 | -13.4046
7.5876 | -34.2199 | -13.4143
7.5937 | -34.2203 | -13.4239
7.5998 | -34.2205 | -13.434
7.6059 | -34.2208 | -13.4436
7.612 | -34.2212 | -13.4537
7.6181 | -34.2214 | -13.4637
7.6242 | -34.2218 | -13.4734
7.6303 | -34.2222 | -13.4828
7.6364 | -34.2226 | -13.493
7.6425 | -34.2229 | -13.5032
7.6486 | -34.2232 | -13.5134
7.6547 | -34.2236 | -13.5236
7.6608 | -34.2239 | -13.5337
7.667 | -34.2243 | -13.543
7.6731 | -34.2246 | -13.5529
7.6792 | -34.2248 | -13.5627
7.6853 | -34.2251 | -13.5724
7.6914 | -34.2254 | -13.5823
7.6975 | -34.2257 | -13.5924
7.7036 | -34.226 | -13.6023
7.7097 | -34.2264 | -13.6119
7.7158 | -34.2267 | -13.6216
7.7219 | -34.2269 | -13.6317
7.728 | -34.2272 | -13.6417
7.7341 | -34.2275 | -13.6515
7.7402 | -34.2279 | -13.6612
7.7463 | -34.2282 | -13.6711
7.7524 | -34.2285 | -13.681
7.7585 | -34.2288 | -13.691
7.7646 | -34.2291 | -13.7007
7.7707 | -34.2294 | -13.7102
7.7768 | -34.2297 | -13.7199
7.7829 | -34.23 | -13.7294
7.789 | -34.2303 | -13.7387
7.7951 | -34.2307 | -13.7489
7.8012 | -34.231 | -13.7592
7.8073 | -34.2312 | -13.7684
7.8135 | -34.2315 | -13.7784
7.8196 | -34.2319 | -13.7879
7.8257 | -34.2322 | -13.7979
7.8318 | -34.2326 | -13.808
7.8379 | -34.2329 | -13.8176
7.844 | -34.2331 | -13.827
7.8501 | -34.2334 | -13.8368
7.8562 | -34.2337 | -13.8469
7.8623 | -34.234 | -13.8568
7.8684 | -34.2343 | -13.8664
7.8745 | -34.2347 | -13.876
7.8806 | -34.235 | -13.885
7.8867 | -34.2353 | -13.8949
7.8928 | -34.2356 | -13.905
7.8989 | -34.236 | -13.9142
7.905 | -34.2363 | -13.9243
7.9111 | -34.2367 | -13.9341
7.9172 | -34.237 | -13.9436
7.9233 | -34.2373 | -13.9533
7.9294 | -34.2377 | -13.963
7.9355 | -34.238 | -13.9726
7.9416 | -34.2383 | -13.9822
7.9477 | -34.2386 | -13.9917
7.9539 | -34.239 | -14.0012
7.96 | -34.2393 | -14.0111
7.9661 | -34.2396 | -14.0209
7.9722 | -34.2399 | -14.031
7.9783 | -34.2402 | -14.0404
7.9844 | -34.2405 | -14.0506
7.9905 | -34.2409 | -14.06
7.9966 | -34.2412 | -14.0696
8.0027 | -34.2415 | -14.0796
8.0088 | -34.2418 | -14.0891
8.0149 | -34.2422 | -14.0985
8.021 | -34.2425 | -14.1081
8.0271 | -34.2428 | -14.1179
8.0332 | -34.2431 | -14.1277
8.0393 | -34.2435 | -14.1373
8.0454 | -34.2438 | -14.1468
8.0515 | -34.2441 | -14.1563
8.0576 | -34.2443 | -14.1659
8.0637 | -34.2447 | -14.1754
8.0698 | -34.245 | -14.185
8.0759 | -34.2454 | -14.1947
8.082 | -34.2457 | -14.2046
8.0881 | -34.246 | -14.2141
8.0942 | -34.2463 | -14.224
8.1004 | -34.2465 | -14.2338
8.1065 | -34.2469 | -14.2436
8.1126 | -34.2473 | -14.2532
8.1187 | -34.2477 | -14.2629
8.1248 | -34.248 | -14.2731
8.1309 | -34.2483 | -14.283
8.137 | -34.2486 | -14.2929
8.1431 | -34.2489 | -14.3024
8.1492 | -34.2493 | -14.3123
8.1553 | -34.2496 | -14.3218
8.1614 | -34.25 | -14.3314
8.1675 | -34.2503 | -14.341
8.1736 | -34.2506 | -14.3508
8.1797 | -34.2509 | -14.3605
8.1858 | -34.2512 | -14.37
8.1919 | -34.2515 | -14.3796
8.198 | -34.2519 | -14.3893
8.2041 | -34.2522 | -14.399
8.2102 | -34.2526 | -14.4087
8.2163 | -34.2529 | -14.4186
8.2224 | -34.2532 | -14.4285
8.2285 | -34.2536 | -14.4384
8.2346 | -34.2539 | -14.4481
8.2408 | -34.2542 | -14.4581
8.2469 | -34.2545 | -14.4675
8.253 | -34.2548 | -14.4772
8.2591 | -34.2552 | -14.4865
8.2652 | -34.2555 | -14.4958
8.2713 | -34.2559 | -14.5054
8.2774 | -34.2562 | -14.515
8.2835 | -34.2566 | -14.5247
8.2896 | -34.2569 | -14.5342
8.2957 | -34.2572 | -14.5437
8.3018 | -34.2576 | -14.5532
8.3079 | -34.258 | -14.5629
8.314 | -34.2583 | -14.5723
8.3201 | -34.2586 | -14.5818
8.3262 | -34.2589 | -14.5913
8.3323 | -34.2593 | -14.601
8.3384 | -34.2596 | -14.6108
8.3445 | -34.26 | -14.6202
8.3506 | -34.2603 | -14.6299
8.3567 | -34.2606 | -14.6397
8.3628 | -34.2609 | -14.6494
8.3689 | -34.2613 | -14.6592
8.375 | -34.2616 | -14.6688
8.3812 | -34.262 | -14.6788
8.3873 | -34.2623 | -14.6885
8.3934 | -34.2627 | -14.698
8.3995 | -34.263 | -14.708
8.4056 | -34.2634 | -14.7177
8.4117 | -34.2637 | -14.7275
8.4178 | -34.2641 | -14.7368
8.4239 | -34.2644 | -14.7465
8.43 | -34.2648 | -14.7562
8.4361 | -34.2652 | -14.7661
8.4422 | -34.2655 | -14.7756
8.4483 | -34.2658 | -14.7853
8.4544 | -34.2661 | -14.7952
8.4605 | -34.2665 | -14.805
8.4666 | -34.2668 | -14.8146
8.4727 | -34.2671 | -14.8241
8.4788 | -34.2675 | -14.8336
8.4849 | -34.2679 | -14.8436
8.491 | -34.2682 | -14.8531
8.4971 | -34.2685 | -14.863
8.5032 | -34.2689 | -14.8729
8.5093 | -34.2692 | -14.8827
8.5154 | -34.2695 | -14.8919
8.5215 | -34.2698 | -14.9018
8.5277 | -34.2702 | -14.9116
8.5338 | -34.2705 | -14.9211
8.5399 | -34.2709 | -14.931
8.546 | -34.2712 | -14.9403
8.5521 | -34.2715 | -14.9501
8.5582 | -34.2719 | -14.96
8.5643 | -34.2722 | -14.9698
8.5704 | -34.2725 | -14.9796
8.5765 | -34.2729 | -14.989
8.5826 | -34.2733 | -14.999
8.5887 | -34.2736 | -15.0086
8.5948 | -34.2739 | -15.0179
8.6009 | -34.2742 | -15.0278
8.607 | -34.2746 | -15.0376
8.6131 | -34.2749 | -15.0474
8.6192 | -34.2752 | -15.0568
8.6253 | -34.2755 | -15.0668
8.6314 | -34.2759 | -15.0766
8.6375 | -34.2762 | -15.0865
8.6436 | -34.2766 | -15.0962
8.6497 | -34.277 | -15.106
8.6558 | -34.2774 | -15.1156
8.6619 | -34.2777 | -15.1254
8.6681 | -34.278 | -15.1347
8.6742 | -34.2783 | -15.1447
8.6803 | -34.2787 | -15.1549
8.6864 | -34.279 | -15.1651
8.6925 | -34.2794 | -15.1751
8.6986 | -34.2797 | -15.185
8.7047 | -34.2801 | -15.1945
8.7108 | -34.2805 | -15.2044
8.7169 | -34.2808 | -15.2137
8.723 | -34.2812 | -15.2235
8.7291 | -34.2815 | -15.2334
8.7352 | -34.2819 | -15.2432
8.7413 | -34.2823 | -15.2527
8.7474 | -34.2826 | -15.2624
8.7535 | -34.283 | -15.2721
8.7596 | -34.2834 | -15.2821
8.7657 | -34.2837 | -15.2917
8.7718 | -34.2841 | -15.3011
8.7779 | -34.2844 | -15.3111
8.784 | -34.2847 | -15.3207
8.7901 | -34.285 | -15.3305
8.7962 | -34.2854 | -15.3407
8.8023 | -34.2858 | -15.3501
8.8084 | -34.2861 | -15.3601
8.8146 | -34.2864 | -15.3697
8.8207 | -34.2868 | -15.3796
8.8268 | -34.2871 | -15.3894
8.8329 | -34.2875 | -15.3993
8.839 | -34.2879 | -15.4089
8.8451 | -34.2882 | -15.4188
8.8512 | -34.2886 | -15.4287
8.8573 | -34.289 | -15.4382
8.8634 | -34.2893 | -15.4481
8.8695 | -34.2897 | -15.4576
8.8756 | -34.29 | -15.4672
8.8817 | -34.2904 | -15.4773
8.8878 | -34.2907 | -15.4869
8.8939 | -34.291 | -15.4964
8.9 | -34.2914 | -15.5057
8.9061 | -34.2917 | -15.5156
8.9122 | -34.2921 | -15.5251
8.9183 | -34.2924 | -15.5346
8.9244 | -34.2927 | -15.5438
8.9305 | -34.2931 | -15.5537
8.9366 | -34.2934 | -15.5632
8.9427 | -34.2939 | -15.5729
8.9488 | -34.2942 | -15.5826
8.955 | -34.2945 | -15.5926
8.9611 | -34.2949 | -15.602
8.9672 | -34.2952 | -15.6119
8.9733 | -34.2956 | -15.6218
8.9794 | -34.2959 | -15.6317
8.9855 | -34.2962 | -15.6412
8.9916 | -34.2966 | -15.6505
8.9977 | -34.297 | -15.6603
9.0038 | -34.2973 | -15.6701
9.0099 | -34.2977 | -15.6798
9.016 | -34.2981 | -15.6895
9.0221 | -34.2984 | -15.6997
9.0282 | -34.2988 | -15.7099
9.0343 | -34.2991 | -15.7198
9.0404 | -34.2995 | -15.7292
9.0465 | -34.2999 | -15.739
9.0526 | -34.3002 | -15.7487
9.0587 | -34.3005 | -15.7583
9.0648 | -34.3009 | -15.768
9.0709 | -34.3013 | -15.7776
9.077 | -34.3016 | -15.7875
9.0831 | -34.3019 | -15.7971
9.0892 | -34.3023 | -15.8068
9.0953 | -34.3025 | -15.8163
9.1015 | -34.3029 | -15.8261
9.1076 | -34.3033 | -15.8359
9.1137 | -34.3037 | -15.8456
9.1198 | -34.304 | -15.855
9.1259 | -34.3044 | -15.8647
9.132 | -34.3048 | -15.8744
9.1381 | -34.3051 | -15.8841
9.1442 | -34.3055 | -15.8936
9.1503 | -34.3058 | -15.9034
9.1564 | -34.3062 | -15.9131
9.1625 | -34.3066 | -15.9228
9.1686 | -34.3069 | -15.9324
9.1747 | -34.3072 | -15.9423
9.1808 | -34.3076 | -15.952
9.1869 | -34.3079 | -15.962
9.193 | -34.3083 | -15.9719
9.1991 | -34.3086 | -15.9813
9.2052 | -34.309 | -15.9909
9.2113 | -34.3094 | -16.0006
9.2174 | -34.3097 | -16.0106
9.2235 | -34.3101 | -16.02
9.2296 | -34.3104 | -16.0298
9.2357 | -34.3108 | -16.0393
9.2419 | -34.3112 | -16.049
9.248 | -34.3115 | -16.0587
9.2541 | -34.3119 | -16.0685
9.2602 | -34.3122 | -16.078
9.2663 | -34.3125 | -16.0878
9.2724 | -34.3129 | -16.0976
9.2785 | -34.3133 | -16.1074
9.2846 | -34.3136 | -16.1171
9.2907 | -34.314 | -16.1268
9.2968 | -34.3143 | -16.1363
9.3029 | -34.3146 | -16.1459
9.309 | -34.315 | -16.1556
9.3151 | -34.3154 | -16.1655
9.3212 | -34.3157 | -16.1753
9.3273 | -34.3161 | -16.1848
9.3334 | -34.3165 | -16.1946
9.3395 | -34.317 | -16.2044
9.3456 | -34.3174 | -16.2142
9.3517 | -34.3178 | -16.2232
9.3578 | -34.3181 | -16.2326
9.3639 | -34.3185 | -16.2426
9.37 | -34.3189 | -16.2522
9.3761 | -34.3192 | -16.2615
9.3822 | -34.3195 | -16.2709
9.3884 | -34.3199 | -16.2808
9.3945 | -34.3203 | -16.2904
9.4006 | -34.3206 | -16.3
9.4067 | -34.3209 | -16.3098
9.4128 | -34.3213 | -16.3192
9.4189 | -34.3217 | -16.3288
9.425 | -34.3221 | -16.3382
9.4311 | -34.3224 | -16.3479
9.4372 | -34.3227 | -16.3575
9.4433 | -34.3231 | -16.3672
9.4494 | -34.3234 | -16.3771
9.4555 | -34.3238 | -16.3865
9.4616 | -34.3242 | -16.3961
9.4677 | -34.3246 | -16.4055
9.4738 | -34.325 | -16.4155
9.4799 | -34.3253 | -16.425
9.486 | -34.3256 | -16.435
9.4921 | -34.326 | -16.445
9.4982 | -34.3263 | -16.4543
9.5043 | -34.3267 | -16.4637
9.5104 | -34.3271 | -16.4733
9.5165 | -34.3274 | -16.4829
9.5226 | -34.3278 | -16.4924
9.5288 | -34.3282 | -16.502
9.5349 | -34.3286 | -16.5116
9.541 | -34.3289 | -16.5214
9.5471 | -34.3293 | -16.5311
9.5532 | -34.3297 | -16.5406
9.5593 | -34.3301 | -16.5499
9.5654 | -34.3305 | -16.5596
9.5715 | -34.3309 | -16.5696
9.5776 | -34.3313 | -16.5793
9.5837 | -34.3317 | -16.5889
9.5898 | -34.332 | -16.5986
9.5959 | -34.3324 | -16.6079
9.602 | -34.3328 | -16.6178
9.6081 | -34.3331 | -16.6276
9.6142 | -34.3335 | -16.6371
9.6203 | -34.3338 | -16.6467
9.6264 | -34.3342 | -16.6563
9.6325 | -34.3345 | -16.6659
9.6386 | -34.335 | -16.6755
9.6447 | -34.3353 | -16.685
9.6508 | -34.3357 | -16.6948
9.6569 | -34.3361 | -16.7044
9.663 | -34.3364 | -16.7139
9.6691 | -34.3369 | -16.7233
9.6753 | -34.3372 | -16.7329
9.6814 | -34.3376 | -16.7425
9.6875 | -34.338 | -16.7522
9.6936 | -34.3384 | -16.7619
9.6997 | -34.3388 | -16.7715
9.7058 | -34.3392 | -16.7808
9.7119 | -34.3395 | -16.7901
9.718 | -34.3398 | -16.7998
9.7241 | -34.3402 | -16.8095
9.7302 | -34.3406 | -16.8192
9.7363 | -34.3409 | -16.8288
9.7424 | -34.3413 | -16.8384
9.7485 | -34.3417 | -16.8479
9.7546 | -34.3421 | -16.8575
9.7607 | -34.3425 | -16.867
9.7668 | -34.3429 | -16.8768
9.7729 | -34.3433 | -16.8862
9.779 | -34.3437 | -16.896
9.7851 | -34.344 | -16.9055
9.7912 | -34.3444 | -16.9156
9.7973 | -34.3447 | -16.9252
9.8034 | -34.3451 | -16.9352
9.8095 | -34.3454 | -16.9443
9.8157 | -34.3458 | -16.9537
9.8218 | -34.3462 | -16.9626
9.8279 | -34.3466 | -16.9725
9.834 | -34.3471 | -16.9818
9.8401 | -34.3474 | -16.9912
9.8462 | -34.3478 | -17.0007
9.8523 | -34.3482 | -17.0102
9.8584 | -34.3486 | -17.0201
9.8645 | -34.349 | -17.0295
9.8706 | -34.3494 | -17.0389
9.8767 | -34.3498 | -17.0482
9.8828 | -34.3503 | -17.0575
9.8889 | -34.3506 | -17.0675
9.895 | -34.351 | -17.0767
9.9011 | -34.3513 | -17.0857
9.9072 | -34.3518 | -17.095
9.9133 | -34.3522 | -17.1045
9.9194 | -34.3526 | -17.1138
9.9255 | -34.353 | -17.1237
9.9316 | -34.3534 | -17.133
9.9377 | -34.3538 | -17.1424
9.9438 | -34.3542 | -17.1518
9.9499 | -34.3546 | -17.1611
9.956 | -34.3551 | -17.1703
9.9622 | -34.3555 | -17.1794
9.9683 | -34.3559 | -17.189
9.9744 | -34.3563 | -17.1982
9.9805 | -34.3567 | -17.2077
9.9866 | -34.3571 | -17.2169
9.9927 | -34.3575 | -17.2269
9.9988 | -34.358 | -17.2359
10.0049 | -34.3583 | -17.2453
10.011 | -34.3588 | -17.2547
10.0171 | -34.3591 | -17.2642
10.0232 | -34.3596 | -17.2735
10.0293 | -34.36 | -17.2826
10.0354 | -34.3604 | -17.2917
10.0415 | -34.3608 | -17.301
10.0476 | -34.3612 | -17.3105
10.0537 | -34.3616 | -17.3202
10.0598 | -34.362 | -17.3298
10.0659 | -34.3624 | -17.3391
10.072 | -34.3628 | -17.3487
10.0781 | -34.3633 | -17.3583
10.0842 | -34.3636 | -17.3678
10.0903 | -34.364 | -17.3776
10.0964 | -34.3645 | -17.387
10.1026 | -34.365 | -17.3964
10.1087 | -34.3655 | -17.4059
10.1148 | -34.3659 | -17.4155
10.1209 | -34.3663 | -17.4249
10.127 | -34.3668 | -17.4342
10.1331 | -34.3672 | -17.4437
10.1392 | -34.3676 | -17.4533
10.1453 | -34.368 | -17.4628
10.1514 | -34.3684 | -17.4719
10.1575 | -34.3688 | -17.4811
10.1636 | -34.3693 | -17.4901
10.1697 | -34.3697 | -17.4994
10.1758 | -34.3701 | -17.5088
10.1819 | -34.3705 | -17.5184
10.188 | -34.3709 | -17.5276
10.1941 | -34.3713 | -17.5371
10.2002 | -34.3717 | -17.546
10.2063 | -34.3722 | -17.5556
10.2124 | -34.3726 | -17.5649
10.2185 | -34.373 | -17.5742
10.2246 | -34.3734 | -17.5836
10.2307 | -34.3738 | -17.593
10.2368 | -34.3743 | -17.602
10.2429 | -34.3747 | -17.6112
10.2491 | -34.3752 | -17.6202
10.2552 | -34.3756 | -17.6299
10.2613 | -34.376 | -17.6394
10.2674 | -34.3764 | -17.6487
10.2735 | -34.3768 | -17.658
10.2796 | -34.3772 | -17.6675
10.2857 | -34.3776 | -17.6767
10.2918 | -34.378 | -17.6859
10.2979 | -34.3785 | -17.6954
10.304 | -34.379 | -17.705
10.3101 | -34.3794 | -17.7144
10.3162 | -34.3799 | -17.7235
10.3223 | -34.3802 | -17.7331
10.3284 | -34.3807 | -17.7424
10.3345 | -34.3811 | -17.7517
10.3406 | -34.3815 | -17.7609
10.3467 | -34.382 | -17.7704
10.3528 | -34.3824 | -17.7797
10.3589 | -34.3828 | -17.7892
10.365 | -34.3832 | -17.7986
10.3711 | -34.3836 | -17.808
10.3772 | -34.384 | -17.8173
10.3833 | -34.3844 | -17.8266
10.3895 | -34.3848 | -17.8357
10.3956 | -34.3853 | -17.8448
10.4017 | -34.3856 | -17.8543
10.4078 | -34.3861 | -17.8637
10.4139 | -34.3866 | -17.8729
10.42 | -34.387 | -17.8821
10.4261 | -34.3875 | -17.8914
10.4322 | -34.3879 | -17.9007
10.4383 | -34.3883 | -17.9101
10.4444 | -34.3888 | -17.9193
10.4505 | -34.3893 | -17.9286
10.4566 | -34.3897 | -17.938
10.4627 | -34.3901 | -17.9476
10.4688 | -34.3906 | -17.9566
10.4749 | -34.391 | -17.9655
10.481 | -34.3914 | -17.9747
10.4871 | -34.3918 | -17.9842
10.4932 | -34.3923 | -17.9936
10.4993 | -34.3928 | -18.0027
10.5054 | -34.3932 | -18.0121
10.5115 | -34.3936 | -18.0214
10.5176 | -34.3941 | -18.0304
10.5237 | -34.3945 | -18.0393
10.5298 | -34.3949 | -18.0486
10.536 | -34.3955 | -18.0581
10.5421 | -34.3958 | -18.0671
10.5482 | -34.3963 | -18.0766
10.5543 | -34.3968 | -18.0859
10.5604 | -34.3972 | -18.0951
10.5665 | -34.3976 | -18.1044
10.5726 | -34.398 | -18.1134
10.5787 | -34.3985 | -18.123
10.5848 | -34.399 | -18.1322
10.5909 | -34.3994 | -18.1414
10.597 | -34.3998 | -18.1506
10.6031 | -34.4003 | -18.1598
10.6092 | -34.4007 | -18.1694
10.6153 | -34.4011 | -18.1787
10.6214 | -34.4016 | -18.188
10.6275 | -34.4021 | -18.1973
10.6336 | -34.4025 | -18.2068
10.6397 | -34.403 | -18.216
10.6458 | -34.4034 | -18.2251
10.6519 | -34.4038 | -18.2342
10.658 | -34.4042 | -18.2432
10.6641 | -34.4047 | -18.2523
10.6702 | -34.4052 | -18.2618
10.6764 | -34.4057 | -18.2708
10.6825 | -34.4061 | -18.2803
10.6886 | -34.4066 | -18.29
10.6947 | -34.407 | -18.2989
10.7008 | -34.4074 | -18.3079
10.7069 | -34.4079 | -18.3171
10.713 | -34.4083 | -18.3262
10.7191 | -34.4088 | -18.3356
10.7252 | -34.4092 | -18.3449
10.7313 | -34.4097 | -18.3543
10.7374 | -34.4102 | -18.3634
10.7435 | -34.4106 | -18.3728
10.7496 | -34.4111 | -18.3821
10.7557 | -34.4115 | -18.3911
10.7618 | -34.412 | -18.4006
10.7679 | -34.4124 | -18.4098
10.774 | -34.4129 | -18.4189
10.7801 | -34.4133 | -18.4285
10.7862 | -34.4138 | -18.4379
10.7923 | -34.4143 | -18.447
10.7984 | -34.4147 | -18.4562
10.8045 | -34.4151 | -18.4658
10.8106 | -34.4156 | -18.475
10.8168 | -34.416 | -18.4845
10.8229 | -34.4165 | -18.4938
10.829 | -34.417 | -18.5033
10.8351 | -34.4174 | -18.5128
10.8412 | -34.4177 | -18.522
10.8473 | -34.4181 | -18.5312
10.8534 | -34.4186 | -18.5406
10.8595 | -34.419 | -18.5497
10.8656 | -34.4195 | -18.559
10.8717 | -34.42 | -18.5679
10.8778 | -34.4204 | -18.5769
10.8839 | -34.4208 | -18.586
10.89 | -34.4213 | -18.5951
10.8961 | -34.4217 | -18.6043
10.9022 | -34.4222 | -18.6135
10.9083 | -34.4227 | -18.6231
10.9144 | -34.4232 | -18.6325
10.9205 | -34.4236 | -18.6416
10.9266 | -34.4241 | -18.6512
10.9327 | -34.4246 | -18.6605
10.9388 | -34.425 | -18.6698
10.9449 | -34.4254 | -18.6791
10.951 | -34.4259 | -18.6882
10.9571 | -34.4263 | -18.6975
10.9633 | -34.4268 | -18.707
10.9694 | -34.4273 | -18.7163
10.9755 | -34.4277 | -18.7253
10.9816 | -34.4282 | -18.7344
10.9877 | -34.4286 | -18.7438
10.9938 | -34.429 | -18.7529
10.9999 | -34.4295 | -18.7621
11.006 | -34.43 | -18.7714
11.0121 | -34.4304 | -18.7811
11.0182 | -34.4309 | -18.79
11.0243 | -34.4313 | -18.7993
11.0304 | -34.4318 | -18.8085
11.0365 | -34.4322 | -18.8177
11.0426 | -34.4326 | -18.8267
11.0487 | -34.4331 | -18.8358
11.0548 | -34.4335 | -18.8451
11.0609 | -34.434 | -18.8545
11.067 | -34.4344 | -18.864
11.0731 | -34.4349 | -18.8733
11.0792 | -34.4353 | -18.8825
11.0853 | -34.4357 | -18.8915
11.0914 | -34.4362 | -18.9004
11.0975 | -34.4367 | -18.9099
11.1037 | -34.4371 | -18.9191
11.1098 | -34.4376 | -18.9282
11.1159 | -34.438 | -18.9378
11.122 | -34.4385 | -18.9475
11.1281 | -34.4389 | -18.9563
11.1342 | -34.4394 | -18.9653
11.1403 | -34.4398 | -18.9746
11.1464 | -34.4403 | -18.9838
11.1525 | -34.4407 | -18.9934
11.1586 | -34.4412 | -19.0029
11.1647 | -34.4417 | -19.0121
11.1708 | -34.4421 | -19.0212
11.1769 | -34.4426 | -19.03
11.183 | -34.443 | -19.0392
11.1891 | -34.4435 | -19.0486
11.1952 | -34.4439 | -19.0582
11.2013 | -34.4443 | -19.0673
11.2074 | -34.4448 | -19.0763
11.2135 | -34.4451 | -19.0853
11.2196 | -34.4456 | -19.0945
11.2257 | -34.446 | -19.1035
11.2318 | -34.4463 | -19.113
11.2379 | -34.4467 | -19.1218
11.244 | -34.4471 | -19.1307
11.2502 | -34.4475 | -19.1398
11.2563 | -34.448 | -19.1489
11.2624 | -34.4484 | -19.1583
11.2685 | -34.4489 | -19.1675
11.2746 | -34.4493 | -19.1766
11.2807 | -34.4497 | -19.1859
11.2868 | -34.4501 | -19.1953
11.2929 | -34.4506 | -19.2044
11.299 | -34.451 | -19.2136
11.3051 | -34.4514 | -19.2225
11.3112 | -34.4518 | -19.2317
11.3173 | -34.4523 | -19.2411
11.3234 | -34.4527 | -19.2505
11.3295 | -34.4531 | -19.2593
11.3356 | -34.4536 | -19.2682
11.3417 | -34.4541 | -19.2778
11.3478 | -34.4545 | -19.2867
11.3539 | -34.4549 | -19.2957
11.36 | -34.4554 | -19.3048
11.3661 | -34.4558 | -19.314
11.3722 | -34.4562 | -19.3233
11.3783 | -34.4567 | -19.3322
11.3844 | -34.4571 | -19.3415
11.3906 | -34.4575 | -19.3509
11.3967 | -34.4579 | -19.3599
11.4028 | -34.4583 | -19.369
11.4089 | -34.4588 | -19.3783
11.415 | -34.4592 | -19.3874
11.4211 | -34.4597 | -19.3964
11.4272 | -34.4601 | -19.4056
11.4333 | -34.4606 | -19.4151
11.4394 | -34.461 | -19.4246
11.4455 | -34.4614 | -19.4341
11.4516 | -34.4619 | -19.4432
11.4577 | -34.4623 | -19.4525
11.4638 | -34.4628 | -19.4614
11.4699 | -34.4632 | -19.4708
11.476 | -34.4637 | -19.4801
11.4821 | -34.4641 | -19.489
11.4882 | -34.4646 | -19.4983
11.4943 | -34.465 | -19.5072
11.5004 | -34.4654 | -19.5165
11.5065 | -34.4659 | -19.5257
11.5126 | -34.4663 | -19.5348
11.5187 | -34.4668 | -19.5439
11.5248 | -34.4672 | -19.5528
11.5309 | -34.4677 | -19.562
11.5371 | -34.4681 | -19.5712
11.5432 | -34.4685 | -19.5805
11.5493 | -34.4689 | -19.5896
11.5554 | -34.4694 | -19.5984
11.5615 | -34.4698 | -19.6076
11.5676 | -34.4703 | -19.6169
11.5737 | -34.4707 | -19.6259
11.5798 | -34.4711 | -19.6351
11.5859 | -34.4716 | -19.644
11.592 | -34.472 | -19.6534
11.5981 | -34.4725 | -19.6626
11.6042 | -34.4729 | -19.6717
11.6103 | -34.4734 | -19.681
11.6164 | -34.4738 | -19.6902
11.6225 | -34.4742 | -19.6994
11.6286 | -34.4747 | -19.7087
11.6347 | -34.4751 | -19.718
11.6408 | -34.4756 | -19.7272
11.6469 | -34.4761 | -19.7366
11.653 | -34.4765 | -19.7456
11.6591 | -34.4769 | -19.7549
11.6652 | -34.4774 | -19.7644
11.6713 | -34.4777 | -19.7736
11.6775 | -34.4782 | -19.7826
11.6836 | -34.4786 | -19.7919
11.6897 | -34.4791 | -19.8008
11.6958 | -34.4795 | -19.8102
11.7019 | -34.48 | -19.8191
11.708 | -34.4804 | -19.8286
11.7141 | -34.4809 | -19.8377
11.7202 | -34.4813 | -19.8467
11.7263 | -34.4818 | -19.856
11.7324 | -34.4823 | -19.8651
11.7385 | -34.4827 | -19.8744
11.7446 | -34.4831 | -19.8839
11.7507 | -34.4835 | -19.8931
11.7568 | -34.484 | -19.9023
11.7629 | -34.4844 | -19.9114
11.769 | -34.4849 | -19.9208
11.7751 | -34.4854 | -19.9297
11.7812 | -34.4858 | -19.9387
11.7873 | -34.4862 | -19.9481
11.7934 | -34.4867 | -19.9572
11.7995 | -34.4871 | -19.9661
11.8056 | -34.4875 | -19.9758
11.8117 | -34.4879 | -19.985
11.8178 | -34.4884 | -19.994
11.824 | -34.4889 | -20.0031
11.8301 | -34.4893 | -20.012
11.8362 | -34.4897 | -20.0211
11.8423 | -34.4901 | -20.0304
11.8484 | -34.4906 | -20.0394
11.8545 | -34.491 | -20.0488
11.8606 | -34.4915 | -20.0582
11.8667 | -34.4919 | -20.0673
11.8728 | -34.4923 | -20.0766
11.8789 | -34.4928 | -20.0862
11.885 | -34.4931 | -20.0951
11.8911 | -34.4937 | -20.1044
11.8972 | -34.4941 | -20.1133
11.9033 | -34.4945 | -20.1222
11.9094 | -34.4949 | -20.1311
11.9155 | -34.4953 | -20.1403
11.9216 | -34.4958 | -20.1492
11.9277 | -34.4962 | -20.1583
11.9338 | -34.4966 | -20.1675
11.9399 | -34.4971 | -20.1767
11.946 | -34.4975 | -20.1859
11.9521 | -34.498 | -20.1951
11.9582 | -34.4984 | -20.204
11.9644 | -34.4989 | -20.2132
11.9705 | -34.4993 | -20.2224
11.9766 | -34.4997 | -20.2315
11.9827 | -34.5001 | -20.2405
11.9888 | -34.5005 | -20.2494
11.9949 | -34.5009 | -20.2588
12.001 | -34.5013 | -20.268
12.0071 | -34.5018 | -20.277
12.0132 | -34.5022 | -20.2861
12.0193 | -34.5027 | -20.2954
12.0254 | -34.5031 | -20.3044
12.0315 | -34.5036 | -20.3135
12.0376 | -34.504 | -20.3225
12.0437 | -34.5044 | -20.3315
12.0498 | -34.5049 | -20.3407
12.0559 | -34.5053 | -20.3501
12.062 | -34.5057 | -20.3593
12.0681 | -34.5061 | -20.3682
12.0742 | -34.5065 | -20.3777
12.0803 | -34.507 | -20.3867
12.0864 | -34.5074 | -20.3956
12.0925 | -34.5079 | -20.4047
12.0986 | -34.5084 | -20.4135
12.1047 | -34.5088 | -20.4227
12.1109 | -34.5092 | -20.4321
12.117 | -34.5097 | -20.4412
12.1231 | -34.5101 | -20.4504
12.1292 | -34.5106 | -20.4597
12.1353 | -34.5111 | -20.4686
12.1414 | -34.5115 | -20.4777
12.1475 | -34.5119 | -20.4867
12.1536 | -34.5123 | -20.4959
12.1597 | -34.5128 | -20.5051
12.1658 | -34.5132 | -20.5143
12.1719 | -34.5137 | -20.5237
12.178 | -34.5142 | -20.5329
12.1841 | -34.5146 | -20.542
12.1902 | -34.515 | -20.5511
12.1963 | -34.5154 | -20.56
12.2024 | -34.5159 | -20.5691
12.2085 | -34.5163 | -20.5782
12.2146 | -34.5168 | -20.5871
12.2207 | -34.5172 | -20.5962
12.2268 | -34.5176 | -20.6053
12.2329 | -34.518 | -20.6144
12.239 | -34.5184 | -20.6233
12.2451 | -34.5189 | -20.6321
12.2513 | -34.5194 | -20.6412
12.2574 | -34.5198 | -20.6505
12.2635 | -34.5202 | -20.6592
12.2696 | -34.5206 | -20.6683
12.2757 | -34.5211 | -20.6772
12.2818 | -34.5215 | -20.686
12.2879 | -34.5219 | -20.6949
12.294 | -34.5223 | -20.7041
12.3001 | -34.5228 | -20.7134
12.3062 | -34.5232 | -20.7224
12.3123 | -34.5237 | -20.7317
12.3184 | -34.5241 | -20.7408
12.3245 | -34.5245 | -20.7502
12.3306 | -34.525 | -20.7591
12.3367 | -34.5254 | -20.7679
12.3428 | -34.5258 | -20.7768
12.3489 | -34.5262 | -20.7855
12.355 | -34.5267 | -20.7944
12.3611 | -34.5271 | -20.8033
12.3672 | -34.5276 | -20.8124
12.3733 | -34.528 | -20.8216
12.3794 | -34.5284 | -20.8305
12.3855 | -34.5288 | -20.8395
12.3916 | -34.5292 | -20.8488
12.3978 | -34.5296 | -20.858
12.4039 | -34.53 | -20.867
12.41 | -34.5305 | -20.8761
12.4161 | -34.5309 | -20.8852
12.4222 | -34.5313 | -20.8941
12.4283 | -34.5318 | -20.9033
12.4344 | -34.5322 | -20.9121
12.4405 | -34.5326 | -20.9211
12.4466 | -34.533 | -20.9302
12.4527 | -34.5335 | -20.9393
12.4588 | -34.5339 | -20.9484
12.4649 | -34.5344 | -20.9576
12.471 | -34.5348 | -20.9669
12.4771 | -34.5352 | -20.9755
12.4832 | -34.5357 | -20.9846
12.4893 | -34.5361 | -20.9935
12.4954 | -34.5366 | -21.0027
12.5015 | -34.5371 | -21.0117
12.5076 | -34.5375 | -21.0206
12.5137 | -34.5378 | -21.03
12.5198 | -34.5383 | -21.0387
12.5259 | -34.5388 | -21.0475
12.532 | -34.5393 | -21.0568
12.5382 | -34.5397 | -21.0656
12.5443 | -34.5401 | -21.0749
12.5504 | -34.5406 | -21.0839
12.5565 | -34.541 | -21.0926
12.5626 | -34.5415 | -21.1016
12.5687 | -34.5419 | -21.1107
12.5748 | -34.5423 | -21.1196
12.5809 | -34.5428 | -21.1283
12.587 | -34.5432 | -21.1375
12.5931 | -34.5436 | -21.1465
12.5992 | -34.544 | -21.1556
12.6053 | -34.5444 | -21.1647
12.6114 | -34.5449 | -21.1739
12.6175 | -34.5453 | -21.1828
12.6236 | -34.5458 | -21.1916
12.6297 | -34.5463 | -21.2003
12.6358 | -34.5467 | -21.2094
12.6419 | -34.5472 | -21.2185
12.648 | -34.5476 | -21.2277
12.6541 | -34.5481 | -21.2366
12.6602 | -34.5485 | -21.2454
12.6663 | -34.5489 | -21.2543
12.6724 | -34.5493 | -21.2634
12.6785 | -34.5498 | -21.2727
12.6847 | -34.5502 | -21.2819
12.6908 | -34.5506 | -21.2908
12.6969 | -34.5511 | -21.2997
12.703 | -34.5516 | -21.3087
12.7091 | -34.552 | -21.3178
12.7152 | -34.5524 | -21.3269
12.7213 | -34.5528 | -21.336
12.7274 | -34.5533 | -21.345
12.7335 | -34.5537 | -21.3537
12.7396 | -34.5541 | -21.3625
12.7457 | -34.5546 | -21.3714
12.7518 | -34.555 | -21.3806
12.7579 | -34.5554 | -21.3895
12.764 | -34.5559 | -21.3985
12.7701 | -34.5563 | -21.407
12.7762 | -34.5567 | -21.4161
12.7823 | -34.5571 | -21.4251
12.7884 | -34.5574 | -21.4338
12.7945 | -34.5579 | -21.4423
12.8006 | -34.5583 | -21.4504
12.8067 | -34.5588 | -21.4597
12.8128 | -34.5592 | -21.4685
12.8189 | -34.5595 | -21.4772
12.8251 | -34.5599 | -21.486
12.8312 | -34.5604 | -21.4946
12.8373 | -34.5608 | -21.5035
12.8434 | -34.5613 | -21.5122
12.8495 | -34.5617 | -21.5212
12.8556 | -34.5621 | -21.5302
12.8617 | -34.5625 | -21.5388
12.8678 | -34.563 | -21.5476
12.8739 | -34.5634 | -21.5562
12.88 | -34.5638 | -21.5653
12.8861 | -34.5642 | -21.5744
12.8922 | -34.5647 | -21.5831
12.8983 | -34.5651 | -21.5915
12.9044 | -34.5656 | -21.6006
12.9105 | -34.5661 | -21.6094
12.9166 | -34.5665 | -21.6186
12.9227 | -34.5669 | -21.6279
12.9288 | -34.5674 | -21.6369
12.9349 | -34.5678 | -21.646
12.941 | -34.5683 | -21.6549
12.9471 | -34.5688 | -21.6636
12.9532 | -34.5692 | -21.6725
12.9593 | -34.5697 | -21.6814
12.9654 | -34.5701 | -21.6902
12.9716 | -34.5705 | -21.6989
12.9777 | -34.5709 | -21.7073
12.9838 | -34.5713 | -21.7161
12.9899 | -34.5718 | -21.7251
12.996 | -34.5722 | -21.7342
13.0021 | -34.5726 | -21.743
13.0082 | -34.573 | -21.7518
13.0143 | -34.5733 | -21.7604
13.0204 | -34.5737 | -21.7692
13.0265 | -34.5742 | -21.7778
13.0326 | -34.5746 | -21.787
13.0387 | -34.5751 | -21.7957
13.0448 | -34.5755 | -21.8046
13.0509 | -34.5759 | -21.8133
13.057 | -34.5764 | -21.822
13.0631 | -34.5768 | -21.8306
13.0692 | -34.5773 | -21.8393
13.0753 | -34.5777 | -21.8479
13.0814 | -34.5782 | -21.8563
13.0875 | -34.5786 | -21.8647
13.0936 | -34.579 | -21.8734
13.0997 | -34.5794 | -21.882
13.1058 | -34.5798 | -21.8902
13.112 | -34.5802 | -21.8989
13.1181 | -34.5807 | -21.9077
13.1242 | -34.5811 | -21.9165
13.1303 | -34.5815 | -21.9254
13.1364 | -34.582 | -21.9342
13.1425 | -34.5824 | -21.943
13.1486 | -34.5829 | -21.9517
13.1547 | -34.5833 | -21.9599
13.1608 | -34.5838 | -21.9687
13.1669 | -34.5842 | -21.9775
13.173 | -34.5846 | -21.9861
13.1791 | -34.585 | -21.9951
13.1852 | -34.5854 | -22.0041
13.1913 | -34.5859 | -22.0129
13.1974 | -34.5863 | -22.0217
13.2035 | -34.5868 | -22.0304
13.2096 | -34.5872 | -22.0391
13.2157 | -34.5876 | -22.0481
13.2218 | -34.5881 | -22.057
13.2279 | -34.5886 | -22.0657
13.234 | -34.589 | -22.0744
13.2401 | -34.5894 | -22.083
13.2462 | -34.5898 | -22.0919
13.2524 | -34.5903 | -22.1003
13.2585 | -34.5907 | -22.109
13.2646 | -34.5911 | -22.1179
13.2707 | -34.5916 | -22.1266
13.2768 | -34.592 | -22.1354
13.2829 | -34.5924 | -22.144
13.289 | -34.5928 | -22.153
13.2951 | -34.5932 | -22.1615
13.3012 | -34.5937 | -22.1703
13.3073 | -34.5941 | -22.1791
13.3134 | -34.5945 | -22.1876
13.3195 | -34.595 | -22.1964
13.3256 | -34.5954 | -22.2054
13.3317 | -34.5959 | -22.2138
13.3378 | -34.5963 | -22.2226
13.3439 | -34.5967 | -22.231
13.35 | -34.5972 | -22.2396
13.3561 | -34.5976 | -22.2486
13.3622 | -34.598 | -22.2575
13.3683 | -34.5985 | -22.2659
13.3744 | -34.5989 | -22.2747
13.3805 | -34.5993 | -22.2834
13.3866 | -34.5997 | -22.292
13.3927 | -34.6001 | -22.3007
13.3989 | -34.6006 | -22.3091
13.405 | -34.6011 | -22.3179
13.4111 | -34.6016 | -22.3267
13.4172 | -34.602 | -22.3351
13.4233 | -34.6024 | -22.3438
13.4294 | -34.6028 | -22.3519
13.4355 | -34.6033 | -22.3613
13.4416 | -34.6037 | -22.37
13.4477 | -34.6041 | -22.3786
13.4538 | -34.6045 | -22.3871
13.4599 | -34.6049 | -22.3959
13.466 | -34.6054 | -22.4048
13.4721 | -34.6059 | -22.4135
13.4782 | -34.6063 | -22.4221
13.4843 | -34.6068 | -22.4305
13.4904 | -34.6072 | -22.4388
13.4965 | -34.6076 | -22.4477
13.5026 | -34.6081 | -22.4563
13.5087 | -34.6086 | -22.4648
13.5148 | -34.609 | -22.4733
13.5209 | -34.6095 | -22.4817
13.527 | -34.6099 | -22.4908
13.5331 | -34.6103 | -22.4993
13.5393 | -34.6107 | -22.5078
13.5454 | -34.6111 | -22.5167
13.5515 | -34.6116 | -22.5251
13.5576 | -34.612 | -22.5339
13.5637 | -34.6125 | -22.5422
13.5698 | -34.6129 | -22.5511
13.5759 | -34.6133 | -22.5596
13.582 | -34.6137 | -22.5684
13.5881 | -34.6141 | -22.5771
13.5942 | -34.6146 | -22.5855
13.6003 | -34.615 | -22.5942
13.6064 | -34.6155 | -22.6029
13.6125 | -34.616 | -22.6116
13.6186 | -34.6164 | -22.6203
13.6247 | -34.6168 | -22.6287
13.6308 | -34.6172 | -22.6368
13.6369 | -34.6177 | -22.6455
13.643 | -34.6182 | -22.6537
13.6491 | -34.6187 | -22.6624
13.6552 | -34.6191 | -22.6707
13.6613 | -34.6195 | -22.6795
13.6674 | -34.62 | -22.6883
13.6735 | -34.6204 | -22.6971
13.6796 | -34.6208 | -22.7055
13.6858 | -34.6213 | -22.7142
13.6919 | -34.6217 | -22.7229
13.698 | -34.6222 | -22.7315
13.7041 | -34.6226 | -22.7401
13.7102 | -34.623 | -22.7489
13.7163 | -34.6235 | -22.7574
13.7224 | -34.6239 | -22.7659
13.7285 | -34.6244 | -22.7745
13.7346 | -34.6248 | -22.783
13.7407 | -34.6252 | -22.7917
13.7468 | -34.6257 | -22.8003
13.7529 | -34.6261 | -22.8089
13.759 | -34.6266 | -22.8175
13.7651 | -34.627 | -22.8261
13.7712 | -34.6274 | -22.8348
13.7773 | -34.6279 | -22.8434
13.7834 | -34.6283 | -22.8519
13.7895 | -34.6288 | -22.8603
13.7956 | -34.6293 | -22.8689
13.8017 | -34.6298 | -22.8777
13.8078 | -34.6302 | -22.8863
13.8139 | -34.6307 | -22.895
13.82 | -34.6312 | -22.9034
13.8262 | -34.6317 | -22.9121
13.8323 | -34.6321 | -22.9207
13.8384 | -34.6326 | -22.9292
13.8445 | -34.633 | -22.9379
13.8506 | -34.6335 | -22.9465
13.8567 | -34.6339 | -22.9551
13.8628 | -34.6344 | -22.9637
13.8689 | -34.6348 | -22.9721
13.875 | -34.6353 | -22.9807
13.8811 | -34.6358 | -22.9893
13.8872 | -34.6363 | -22.998
13.8933 | -34.6367 | -23.0064
13.8994 | -34.6372 | -23.0147
13.9055 | -34.6377 | -23.0229
13.9116 | -34.6381 | -23.0313
13.9177 | -34.6386 | -23.0395
13.9238 | -34.6391 | -23.0482
13.9299 | -34.6396 | -23.0567
13.936 | -34.6401 | -23.0651
13.9421 | -34.6405 | -23.0738
13.9482 | -34.641 | -23.0825
13.9543 | -34.6415 | -23.0907
13.9604 | -34.642 | -23.0995
13.9665 | -34.6425 | -23.1079
13.9727 | -34.6429 | -23.1161
13.9788 | -34.6433 | -23.1248
13.9849 | -34.6437 | -23.1334
13.991 | -34.6441 | -23.1422
13.9971 | -34.6446 | -23.1511
14.0032 | -34.6451 | -23.1598
14.0093 | -34.6454 | -23.1681
14.0154 | -34.6459 | -23.177
14.0215 | -34.6463 | -23.1854
14.0276 | -34.6468 | -23.1937
14.0337 | -34.6473 | -23.2023
14.0398 | -34.6477 | -23.2107
14.0459 | -34.6481 | -23.2194
14.052 | -34.6485 | -23.2279
14.0581 | -34.649 | -23.2361
14.0642 | -34.6495 | -23.2449
14.0703 | -34.65 | -23.2533
14.0764 | -34.6504 | -23.2624
14.0825 | -34.6509 | -23.2709
14.0886 | -34.6513 | -23.2789
14.0947 | -34.6517 | -23.2871
14.1008 | -34.6522 | -23.296
14.1069 | -34.6527 | -23.3043
14.1131 | -34.6531 | -23.313
14.1192 | -34.6536 | -23.3212
14.1253 | -34.654 | -23.3296
14.1314 | -34.6545 | -23.3381
14.1375 | -34.655 | -23.3465
14.1436 | -34.6554 | -23.355
14.1497 | -34.6559 | -23.3633
14.1558 | -34.6564 | -23.372
14.1619 | -34.6568 | -23.3806
14.168 | -34.6573 | -23.389
14.1741 | -34.6578 | -23.3972
14.1802 | -34.6582 | -23.4057
14.1863 | -34.6587 | -23.4142
14.1924 | -34.6591 | -23.4227
14.1985 | -34.6595 | -23.4309
14.2046 | -34.66 | -23.439
14.2107 | -34.6605 | -23.4471
14.2168 | -34.661 | -23.4559
14.2229 | -34.6615 | -23.4644
14.229 | -34.6619 | -23.4727
14.2351 | -34.6624 | -23.481
14.2412 | -34.6629 | -23.4893
14.2473 | -34.6634 | -23.4977
14.2534 | -34.6639 | -23.5063
14.2596 | -34.6643 | -23.5149
14.2657 | -34.6648 | -23.5236
14.2718 | -34.6653 | -23.5321
14.2779 | -34.6657 | -23.5402
14.284 | -34.6662 | -23.5483
14.2901 | -34.6667 | -23.5566
14.2962 | -34.6672 | -23.5646
14.3023 | -34.6676 | -23.5731
14.3084 | -34.6681 | -23.5819
14.3145 | -34.6686 | -23.59
14.3206 | -34.6691 | -23.5982
14.3267 | -34.6695 | -23.6065
14.3328 | -34.67 | -23.6148
14.3389 | -34.6704 | -23.623
14.345 | -34.6709 | -23.6316
14.3511 | -34.6714 | -23.6397
14.3572 | -34.6718 | -23.6482
14.3633 | -34.6723 | -23.6563
14.3694 | -34.6727 | -23.6647
14.3755 | -34.6731 | -23.6727
14.3816 | -34.6736 | -23.6807
14.3877 | -34.6741 | -23.6889
14.3938 | -34.6746 | -23.6973
14.4 | -34.675 | -23.7056
14.4061 | -34.6755 | -23.714
14.4122 | -34.6759 | -23.7225
14.4183 | -34.6764 | -23.7311
14.4244 | -34.6768 | -23.7391
14.4305 | -34.6773 | -23.7477
14.4366 | -34.6777 | -23.7561
14.4427 | -34.6782 | -23.7645
14.4488 | -34.6787 | -23.7725
14.4549 | -34.6792 | -23.7806
14.461 | -34.6796 | -23.7892
14.4671 | -34.6801 | -23.7976
14.4732 | -34.6806 | -23.8055
14.4793 | -34.6811 | -23.8145
14.4854 | -34.6816 | -23.8224
14.4915 | -34.6821 | -23.8307
14.4976 | -34.6826 | -23.8393
14.5037 | -34.683 | -23.8478
14.5098 | -34.6835 | -23.8559
14.5159 | -34.684 | -23.8645
14.522 | -34.6845 | -23.8731
14.5281 | -34.685 | -23.8811
14.5342 | -34.6854 | -23.8895
14.5403 | -34.6859 | -23.898
14.5465 | -34.6864 | -23.9066
14.5526 | -34.6869 | -23.9151
14.5587 | -34.6873 | -23.9233
14.5648 | -34.6878 | -23.9314
14.5709 | -34.6883 | -23.94
14.577 | -34.6887 | -23.9485
14.5831 | -34.6892 | -23.9569
14.5892 | -34.6898 | -23.9648
14.5953 | -34.6902 | -23.9728
14.6014 | -34.6907 | -23.9811
14.6075 | -34.6913 | -23.9893
14.6136 | -34.6917 | -23.9975
14.6197 | -34.6922 | -24.0059
14.6258 | -34.6927 | -24.0142
14.6319 | -34.6932 | -24.0222
14.638 | -34.6938 | -24.0306
14.6441 | -34.6942 | -24.0384
14.6502 | -34.6947 | -24.0462
14.6563 | -34.6952 | -24.0543
14.6624 | -34.6956 | -24.0625
14.6685 | -34.6961 | -24.0704
14.6746 | -34.6966 | -24.0787
14.6807 | -34.6971 | -24.0872
14.6869 | -34.6975 | -24.0954
14.693 | -34.698 | -24.1036
14.6991 | -34.6984 | -24.1119
14.7052 | -34.6989 | -24.1199
14.7113 | -34.6993 | -24.1281
14.7174 | -34.6998 | -24.1362
14.7235 | -34.7003 | -24.1444
14.7296 | -34.7008 | -24.1527
14.7357 | -34.7013 | -24.1605
14.7418 | -34.7018 | -24.1685
14.7479 | -34.7022 | -24.1769
14.754 | -34.7027 | -24.1849
14.7601 | -34.7032 | -24.1934
14.7662 | -34.7037 | -24.2018
14.7723 | -34.7042 | -24.21
14.7784 | -34.7047 | -24.2183
14.7845 | -34.7052 | -24.2266
14.7906 | -34.7057 | -24.2347
14.7967 | -34.7061 | -24.2429
14.8028 | -34.7066 | -24.2511
14.8089 | -34.7071 | -24.259
14.815 | -34.7075 | -24.2669
14.8211 | -34.708 | -24.275
14.8272 | -34.7085 | -24.2832
14.8334 | -34.7089 | -24.2913
14.8395 | -34.7094 | -24.2995
14.8456 | -34.7099 | -24.3076
14.8517 | -34.7104 | -24.3158
14.8578 | -34.7109 | -24.3238
14.8639 | -34.7114 | -24.332
14.87 | -34.7119 | -24.34
14.8761 | -34.7124 | -24.3477
14.8822 | -34.7128 | -24.3558
14.8883 | -34.7133 | -24.3638
14.8944 | -34.7139 | -24.3721
14.9005 | -34.7143 | -24.38
14.9066 | -34.7148 | -24.3882
14.9127 | -34.7153 | -24.3961
14.9188 | -34.7158 | -24.4045
14.9249 | -34.7163 | -24.4126
14.931 | -34.7167 | -24.4204
14.9371 | -34.7172 | -24.4284
14.9432 | -34.7177 | -24.4365
14.9493 | -34.7182 | -24.4445
14.9554 | -34.7187 | -24.4526
14.9615 | -34.7192 | -24.4607
14.9676 | -34.7197 | -24.4688
14.9738 | -34.7201 | -24.4769
14.9799 | -34.7206 | -24.4851
14.986 | -34.7211 | -24.4928
14.9921 | -34.7217 | -24.5008
14.9982 | -34.7221 | -24.5093
15.0043 | -34.7226 | -24.5175
15.0104 | -34.7231 | -24.5254
15.0165 | -34.7236 | -24.5334
15.0226 | -34.7241 | -24.5411
15.0287 | -34.7245 | -24.5496
15.0348 | -34.725 | -24.5577
15.0409 | -34.7255 | -24.5659
15.047 | -34.726 | -24.574
15.0531 | -34.7266 | -24.5821
15.0592 | -34.7271 | -24.5904
15.0653 | -34.7276 | -24.5982
15.0714 | -34.7281 | -24.6063
15.0775 | -34.7285 | -24.6142
15.0836 | -34.729 | -24.6222
15.0897 | -34.7295 | -24.6302
15.0958 | -34.73 | -24.6382
15.1019 | -34.7304 | -24.6461
15.108 | -34.7309 | -24.6543
15.1141 | -34.7313 | -24.6625
15.1203 | -34.7319 | -24.6704
15.1264 | -34.7324 | -24.6784
15.1325 | -34.7329 | -24.6863
15.1386 | -34.7334 | -24.6945
15.1447 | -34.7339 | -24.7025
15.1508 | -34.7344 | -24.7107
15.1569 | -34.7349 | -24.7189
15.163 | -34.7355 | -24.7266
15.1691 | -34.7359 | -24.7345
15.1752 | -34.7364 | -24.7426
15.1813 | -34.7369 | -24.7507
15.1874 | -34.7375 | -24.7589
15.1935 | -34.738 | -24.7667
15.1996 | -34.7385 | -24.7751
15.2057 | -34.739 | -24.7831
15.2118 | -34.7395 | -24.7912
15.2179 | -34.74 | -24.7988
15.224 | -34.7404 | -24.8072
15.2301 | -34.741 | -24.815
15.2362 | -34.7415 | -24.823
15.2423 | -34.742 | -24.8312
15.2484 | -34.7424 | -24.839
15.2545 | -34.743 | -24.8469
15.2607 | -34.7434 | -24.855
15.2668 | -34.744 | -24.8628
15.2729 | -34.7445 | -24.8708
15.279 | -34.745 | -24.8787
15.2851 | -34.7454 | -24.8868
15.2912 | -34.7459 | -24.8945
15.2973 | -34.7464 | -24.9022
15.3034 | -34.747 | -24.9099
15.3095 | -34.7474 | -24.9176
15.3156 | -34.7478 | -24.9258
15.3217 | -34.7483 | -24.9335
15.3278 | -34.7489 | -24.9414
15.3339 | -34.7494 | -24.949
15.34 | -34.7498 | -24.9571
15.3461 | -34.7503 | -24.9651
15.3522 | -34.7508 | -24.9731
15.3583 | -34.7513 | -24.9808
15.3644 | -34.7518 | -24.989
15.3705 | -34.7523 | -24.9969
15.3766 | -34.7528 | -25.0049
15.3827 | -34.7533 | -25.0128
15.3888 | -34.7537 | -25.0203
15.3949 | -34.7542 | -25.0281
15.401 | -34.7548 | -25.0361
15.4072 | -34.7553 | -25.0438
15.4133 | -34.7558 | -25.0518
15.4194 | -34.7563 | -25.0598
15.4255 | -34.7568 | -25.0676
15.4316 | -34.7574 | -25.0756
15.4377 | -34.7579 | -25.0832
15.4438 | -34.7584 | -25.0905
15.4499 | -34.7588 | -25.0985
15.456 | -34.7592 | -25.1065
15.4621 | -34.7598 | -25.1145
15.4682 | -34.7603 | -25.1224
15.4743 | -34.7609 | -25.1301
15.4804 | -34.7614 | -25.1381
15.4865 | -34.7619 | -25.1461
15.4926 | -34.7623 | -25.154
15.4987 | -34.7628 | -25.162
15.5048 | -34.7633 | -25.1698
15.5109 | -34.7639 | -25.1773
15.517 | -34.7644 | -25.1852
15.5231 | -34.7649 | -25.1931
15.5292 | -34.7654 | -25.2011
15.5353 | -34.7659 | -25.2088
15.5414 | -34.7664 | -25.2169
15.5476 | -34.7669 | -25.2249
15.5537 | -34.7674 | -25.2327
15.5598 | -34.7679 | -25.2402
15.5659 | -34.7684 | -25.2484
15.572 | -34.7689 | -25.2558
15.5781 | -34.7694 | -25.264
15.5842 | -34.7699 | -25.2716
15.5903 | -34.7704 | -25.2794
15.5964 | -34.7709 | -25.2872
15.6025 | -34.7714 | -25.295
15.6086 | -34.7718 | -25.3029
15.6147 | -34.7723 | -25.3107
15.6208 | -34.7729 | -25.3184
15.6269 | -34.7734 | -25.3263
15.633 | -34.7739 | -25.3341
15.6391 | -34.7743 | -25.3419
15.6452 | -34.7749 | -25.3495
15.6513 | -34.7753 | -25.3577
15.6574 | -34.7758 | -25.3651
15.6635 | -34.7763 | -25.3729
15.6696 | -34.7768 | -25.3805
15.6757 | -34.7773 | -25.3886
15.6818 | -34.7778 | -25.3968
15.688 | -34.7784 | -25.4046
15.6941 | -34.7789 | -25.4126
15.7002 | -34.7794 | -25.4206
15.7063 | -34.7799 | -25.4286
15.7124 | -34.7804 | -25.4365
15.7185 | -34.7809 | -25.4443
15.7246 | -34.7814 | -25.4517
15.7307 | -34.7819 | -25.4597
15.7368 | -34.7824 | -25.4678
15.7429 | -34.7829 | -25.4755
15.749 | -34.7834 | -25.4831
15.7551 | -34.7839 | -25.4908
15.7612 | -34.7843 | -25.4978
15.7673 | -34.7848 | -25.5056
15.7734 | -34.7853 | -25.5132
15.7795 | -34.7858 | -25.5214
15.7856 | -34.7863 | -25.5291
15.7917 | -34.7867 | -25.5368
15.7978 | -34.7873 | -25.5444
15.8039 | -34.7877 | -25.5521
15.81 | -34.7882 | -25.5598
15.8161 | -34.7887 | -25.5679
15.8222 | -34.7892 | -25.5759
15.8283 | -34.7897 | -25.5835
15.8345 | -34.7902 | -25.5913
15.8406 | -34.7907 | -25.5992
15.8467 | -34.7912 | -25.6068
15.8528 | -34.7917 | -25.6145
15.8589 | -34.7922 | -25.6222
15.865 | -34.7926 | -25.6299
15.8711 | -34.7931 | -25.6379
15.8772 | -34.7937 | -25.6458
15.8833 | -34.7942 | -25.6538
15.8894 | -34.7947 | -25.6614
15.8955 | -34.7952 | -25.6695
15.9016 | -34.7957 | -25.6772
15.9077 | -34.7962 | -25.6851
15.9138 | -34.7967 | -25.6928
15.9199 | -34.7972 | -25.7008
15.926 | -34.7977 | -25.709
15.9321 | -34.7982 | -25.7167
15.9382 | -34.7988 | -25.7245
15.9443 | -34.7993 | -25.7325
15.9504 | -34.7998 | -25.7401
15.9565 | -34.8004 | -25.7478
15.9626 | -34.8009 | -25.7556
15.9687 | -34.8014 | -25.7633
15.9749 | -34.802 | -25.7714
15.981 | -34.8025 | -25.7792
15.9871 | -34.803 | -25.7869
15.9932 | -34.8035 | -25.7949
15.9993 | -34.8039 | -25.8024
16.0054 | -34.8044 | -25.8101
16.0115 | -34.805 | -25.8176
16.0176 | -34.8055 | -25.8254
16.0237 | -34.806 | -25.8334
16.0298 | -34.8065 | -25.841
16.0359 | -34.807 | -25.8487
16.042 | -34.8076 | -25.8563
16.0481 | -34.8081 | -25.8641
16.0542 | -34.8085 | -25.872
16.0603 | -34.8091 | -25.8798
16.0664 | -34.8096 | -25.8877
16.0725 | -34.8101 | -25.8954
16.0786 | -34.8106 | -25.9035
16.0847 | -34.8111 | -25.9108
16.0908 | -34.8116 | -25.9184
16.0969 | -34.8122 | -25.9258
16.103 | -34.8126 | -25.9338
16.1091 | -34.8131 | -25.9412
16.1152 | -34.8136 | -25.9492
16.1214 | -34.8141 | -25.9569
16.1275 | -34.8146 | -25.9646
16.1336 | -34.8151 | -25.9723
16.1397 | -34.8157 | -25.98
16.1458 | -34.8162 | -25.9877
16.1519 | -34.8167 | -25.9955
16.158 | -34.8172 | -26.0033
16.1641 | -34.8177 | -26.0111
16.1702 | -34.8182 | -26.0189
16.1763 | -34.8187 | -26.0266
16.1824 | -34.8192 | -26.0342
16.1885 | -34.8198 | -26.0421
16.1946 | -34.8202 | -26.0498
16.2007 | -34.8207 | -26.0579
16.2068 | -34.8212 | -26.0653
16.2129 | -34.8217 | -26.073
16.219 | -34.8221 | -26.0809
16.2251 | -34.8226 | -26.0889
16.2312 | -34.823 | -26.0967
16.2373 | -34.8235 | -26.1046
16.2434 | -34.824 | -26.1123
16.2495 | -34.8245 | -26.1198
16.2556 | -34.825 | -26.1274
16.2618 | -34.8255 | -26.1348
16.2679 | -34.826 | -26.1424
16.274 | -34.8265 | -26.1503
16.2801 | -34.827 | -26.1578
16.2862 | -34.8275 | -26.1653
16.2923 | -34.828 | -26.1732
16.2984 | -34.8285 | -26.1811
16.3045 | -34.8291 | -26.1888
16.3106 | -34.8296 | -26.1962
16.3167 | -34.8301 | -26.2041
16.3228 | -34.8306 | -26.2118
16.3289 | -34.8311 | -26.2196
16.335 | -34.8316 | -26.2276
16.3411 | -34.8321 | -26.2352
16.3472 | -34.8326 | -26.2429
16.3533 | -34.833 | -26.2507
16.3594 | -34.8335 | -26.2585
16.3655 | -34.834 | -26.2664
16.3716 | -34.8345 | -26.2742
16.3777 | -34.835 | -26.2818
16.3838 | -34.8355 | -26.2895
16.3899 | -34.836 | -26.2972
16.396 | -34.8365 | -26.305
16.4021 | -34.837 | -26.3128
16.4083 | -34.8375 | -26.3206
16.4144 | -34.838 | -26.3281
16.4205 | -34.8385 | -26.3359
16.4266 | -34.839 | -26.3437
16.4327 | -34.8395 | -26.3513
16.4388 | -34.84 | -26.3592
16.4449 | -34.8405 | -26.3669
16.451 | -34.841 | -26.3748
16.4571 | -34.8415 | -26.3825
16.4632 | -34.842 | -26.3902
16.4693 | -34.8425 | -26.3981
16.4754 | -34.843 | -26.4057
16.4815 | -34.8435 | -26.4135
16.4876 | -34.844 | -26.4214
16.4937 | -34.8445 | -26.4292
16.4998 | -34.845 | -26.4368
16.5059 | -34.8455 | -26.4444
16.512 | -34.846 | -26.4523
16.5181 | -34.8465 | -26.4599
16.5242 | -34.8471 | -26.4676
16.5303 | -34.8476 | -26.4751
16.5364 | -34.8481 | -26.4828
16.5425 | -34.8486 | -26.4906
16.5487 | -34.8491 | -26.4984
16.5548 | -34.8496 | -26.5061
16.5609 | -34.8501 | -26.5139
16.567 | -34.8505 | -26.5216
16.5731 | -34.851 | -26.5292
16.5792 | -34.8515 | -26.5368
16.5853 | -34.8521 | -26.5443
16.5914 | -34.8526 | -26.5521
16.5975 | -34.8531 | -26.5599
16.6036 | -34.8536 | -26.5676
16.6097 | -34.8541 | -26.5753
16.6158 | -34.8546 | -26.5833
16.6219 | -34.8551 | -26.5907
16.628 | -34.8556 | -26.5985
16.6341 | -34.8561 | -26.6059
16.6402 | -34.8566 | -26.6133
16.6463 | -34.8572 | -26.621
16.6524 | -34.8576 | -26.6287
16.6585 | -34.8581 | -26.6364
16.6646 | -34.8586 | -26.6441
16.6707 | -34.8591 | -26.6516
16.6768 | -34.8596 | -26.6593
16.6829 | -34.8601 | -26.6671
16.689 | -34.8606 | -26.6748
16.6952 | -34.8611 | -26.6826
16.7013 | -34.8616 | -26.6903
16.7074 | -34.8621 | -26.6981
16.7135 | -34.8626 | -26.7058
16.7196 | -34.8631 | -26.7136
16.7257 | -34.8635 | -26.7211
16.7318 | -34.864 | -26.7288
16.7379 | -34.8645 | -26.7368
16.744 | -34.865 | -26.7442
16.7501 | -34.8655 | -26.7519
16.7562 | -34.866 | -26.7596
16.7623 | -34.8665 | -26.7675
16.7684 | -34.867 | -26.7753
16.7745 | -34.8675 | -26.7826
16.7806 | -34.868 | -26.7905
16.7867 | -34.8685 | -26.798
16.7928 | -34.8689 | -26.8058
16.7989 | -34.8694 | -26.8136
16.805 | -34.8699 | -26.8213
16.8111 | -34.8704 | -26.8289
16.8172 | -34.8709 | -26.8368
16.8233 | -34.8714 | -26.8446
16.8294 | -34.8719 | -26.8524
16.8356 | -34.8724 | -26.8602
16.8417 | -34.8729 | -26.8679
16.8478 | -34.8734 | -26.8754
16.8539 | -34.8739 | -26.8832
16.86 | -34.8744 | -26.8906
16.8661 | -34.8749 | -26.8983
16.8722 | -34.8754 | -26.9058
16.8783 | -34.8759 | -26.9129
16.8844 | -34.8763 | -26.9206
16.8905 | -34.8769 | -26.9284
16.8966 | -34.8773 | -26.9361
16.9027 | -34.8778 | -26.9435
16.9088 | -34.8783 | -26.9515
16.9149 | -34.8789 | -26.9593
16.921 | -34.8793 | -26.9669
16.9271 | -34.8799 | -26.9747
16.9332 | -34.8804 | -26.9825
16.9393 | -34.8808 | -26.9904
16.9454 | -34.8813 | -26.998
16.9515 | -34.8818 | -27.0056
16.9576 | -34.8823 | -27.0133
16.9637 | -34.8828 | -27.0212
16.9698 | -34.8833 | -27.0288
16.9759 | -34.8838 | -27.0366
16.9821 | -34.8843 | -27.0443
16.9882 | -34.8847 | -27.0521
16.9943 | -34.8852 | -27.0601
17.0004 | -34.8857 | -27.0679
17.0065 | -34.8862 | -27.0755
17.0126 | -34.8867 | -27.0831
17.0187 | -34.8872 | -27.0907
17.0248 | -34.8877 | -27.0979
17.0309 | -34.8882 | -27.1057
17.037 | -34.8886 | -27.1136
17.0431 | -34.8892 | -27.1209
17.0492 | -34.8896 | -27.1287
17.0553 | -34.8901 | -27.1363
17.0614 | -34.8906 | -27.1437
17.0675 | -34.8911 | -27.1516
17.0736 | -34.8916 | -27.1593
17.0797 | -34.8921 | -27.1667
17.0858 | -34.8926 | -27.1744
17.0919 | -34.893 | -27.1822
17.098 | -34.8935 | -27.1898
17.1041 | -34.894 | -27.1974
17.1102 | -34.8945 | -27.2052
17.1163 | -34.8949 | -27.213
17.1225 | -34.8954 | -27.2207
17.1286 | -34.8959 | -27.2282
17.1347 | -34.8964 | -27.2357
17.1408 | -34.8969 | -27.2431
17.1469 | -34.8974 | -27.2506
17.153 | -34.8979 | -27.2582
17.1591 | -34.8984 | -27.2657
17.1652 | -34.8989 | -27.2734
17.1713 | -34.8994 | -27.2811
17.1774 | -34.8999 | -27.289
17.1835 | -34.9004 | -27.2968
17.1896 | -34.9009 | -27.3044
17.1957 | -34.9013 | -27.3121
17.2018 | -34.9018 | -27.32
17.2079 | -34.9022 | -27.3273
17.214 | -34.9027 | -27.3349
17.2201 | -34.9032 | -27.3426
17.2262 | -34.9037 | -27.3503
17.2323 | -34.9041 | -27.3578
17.2384 | -34.9047 | -27.365
17.2445 | -34.9051 | -27.3729
17.2506 | -34.9056 | -27.3806
17.2567 | -34.906 | -27.3883
17.2628 | -34.9065 | -27.3962
17.269 | -34.907 | -27.4036
17.2751 | -34.9075 | -27.4112
17.2812 | -34.908 | -27.4187
17.2873 | -34.9085 | -27.4264
17.2934 | -34.9089 | -27.4338
17.2995 | -34.9094 | -27.4412
17.3056 | -34.9098 | -27.4488
17.3117 | -34.9103 | -27.4563
17.3178 | -34.9108 | -27.464
17.3239 | -34.9113 | -27.4715
17.33 | -34.9118 | -27.4792
17.3361 | -34.9122 | -27.4867
17.3422 | -34.9127 | -27.4944
17.3483 | -34.9132 | -27.5021
17.3544 | -34.9137 | -27.5094
17.3605 | -34.9142 | -27.5173
17.3666 | -34.9147 | -27.525
17.3727 | -34.9152 | -27.5325
17.3788 | -34.9156 | -27.54
17.3849 | -34.9161 | -27.5476
17.391 | -34.9166 | -27.555
17.3971 | -34.9171 | -27.5627
17.4032 | -34.9176 | -27.5707
17.4094 | -34.9181 | -27.5782
17.4155 | -34.9185 | -27.5858
17.4216 | -34.919 | -27.5935
17.4277 | -34.9195 | -27.6011
17.4338 | -34.92 | -27.6089
17.4399 | -34.9204 | -27.6164
17.446 | -34.9209 | -27.6239
17.4521 | -34.9214 | -27.6313
17.4582 | -34.9218 | -27.639
17.4643 | -34.9223 | -27.6468
17.4704 | -34.9227 | -27.6542
17.4765 | -34.9232 | -27.6617
17.4826 | -34.9237 | -27.669
17.4887 | -34.9242 | -27.6768
17.4948 | -34.9247 | -27.6843
17.5009 | -34.9252 | -27.6917
17.507 | -34.9257 | -27.6994
17.5131 | -34.9261 | -27.7069
17.5192 | -34.9266 | -27.7145
17.5253 | -34.927 | -27.7221
17.5314 | -34.9275 | -27.7298
17.5375 | -34.928 | -27.7374
17.5436 | -34.9284 | -27.745
17.5497 | -34.9289 | -27.7528
17.5559 | -34.9294 | -27.7602
17.562 | -34.9298 | -27.7677
17.5681 | -34.9303 | -27.7752
17.5742 | -34.9308 | -27.7828
17.5803 | -34.9312 | -27.7902
17.5864 | -34.9317 | -27.7972
17.5925 | -34.9322 | -27.8048
17.5986 | -34.9327 | -27.8121
17.6047 | -34.9332 | -27.8199
17.6108 | -34.9337 | -27.8274
17.6169 | -34.9342 | -27.8349
17.623 | -34.9347 | -27.8423
17.6291 | -34.9351 | -27.85
17.6352 | -34.9356 | -27.8574
17.6413 | -34.936 | -27.865
17.6474 | -34.9365 | -27.8725
17.6535 | -34.937 | -27.8802
17.6596 | -34.9375 | -27.8876
17.6657 | -34.938 | -27.8949
17.6718 | -34.9384 | -27.9027
17.6779 | -34.9389 | -27.9103
17.684 | -34.9394 | -27.9178
17.6901 | -34.9399 | -27.9252
17.6963 | -34.9403 | -27.9325
17.7024 | -34.9408 | -27.9401
17.7085 | -34.9412 | -27.9475
17.7146 | -34.9417 | -27.955
17.7207 | -34.9422 | -27.9626
17.7268 | -34.9427 | -27.9699
17.7329 | -34.9431 | -27.9777
17.739 | -34.9436 | -27.9853
17.7451 | -34.9441 | -27.9928
17.7512 | -34.9446 | -28.0004
17.7573 | -34.945 | -28.0077
17.7634 | -34.9455 | -28.015
17.7695 | -34.946 | -28.0226
17.7756 | -34.9465 | -28.0301
17.7817 | -34.947 | -28.0374
17.7878 | -34.9475 | -28.0449
17.7939 | -34.9479 | -28.0523
17.8 | -34.9484 | -28.0597
17.8061 | -34.9489 | -28.0672
17.8122 | -34.9493 | -28.0747
17.8183 | -34.9498 | -28.0824
17.8244 | -34.9503 | -28.0898
17.8305 | -34.9508 | -28.0973
17.8366 | -34.9512 | -28.1045
17.8428 | -34.9517 | -28.1122
17.8489 | -34.9522 | -28.1196
17.855 | -34.9526 | -28.127
17.8611 | -34.9531 | -28.1345
17.8672 | -34.9536 | -28.1418
17.8733 | -34.9541 | -28.1493
17.8794 | -34.9545 | -28.1569
17.8855 | -34.955 | -28.1643
17.8916 | -34.9555 | -28.1716
17.8977 | -34.956 | -28.179
17.9038 | -34.9565 | -28.1861
17.9099 | -34.9569 | -28.1936
17.916 | -34.9574 | -28.2007
17.9221 | -34.9579 | -28.2082
17.9282 | -34.9583 | -28.2156
17.9343 | -34.9588 | -28.2231
17.9404 | -34.9593 | -28.2305
17.9465 | -34.9598 | -28.238
17.9526 | -34.9602 | -28.2455
17.9587 | -34.9607 | -28.2529
17.9648 | -34.9612 | -28.2605
17.9709 | -34.9617 | -28.268
17.977 | -34.9622 | -28.275
17.9832 | -34.9627 | -28.2825
17.9893 | -34.9631 | -28.29
17.9954 | -34.9636 | -28.2977
18.0015 | -34.9641 | -28.3051
18.0076 | -34.9646 | -28.3119
18.0137 | -34.965 | -28.3197
18.0198 | -34.9655 | -28.3273
18.0259 | -34.9659 | -28.3347
18.032 | -34.9664 | -28.3419
18.0381 | -34.9668 | -28.3492
18.0442 | -34.9674 | -28.3567
18.0503 | -34.9678 | -28.3642
18.0564 | -34.9683 | -28.3716
18.0625 | -34.9687 | -28.3789
18.0686 | -34.9692 | -28.3862
18.0747 | -34.9697 | -28.3935
18.0808 | -34.9702 | -28.4009
18.0869 | -34.9707 | -28.4084
18.093 | -34.9712 | -28.4156
18.0991 | -34.9717 | -28.4231
18.1052 | -34.9722 | -28.4307
18.1113 | -34.9726 | -28.4378
18.1174 | -34.9731 | -28.4449
18.1236 | -34.9736 | -28.4523
18.1297 | -34.9741 | -28.4598
18.1358 | -34.9746 | -28.4672
18.1419 | -34.9751 | -28.4747
18.148 | -34.9756 | -28.4823
18.1541 | -34.9761 | -28.4894
18.1602 | -34.9766 | -28.4968
18.1663 | -34.9771 | -28.5042
18.1724 | -34.9775 | -28.5118
18.1785 | -34.978 | -28.5194
18.1846 | -34.9785 | -28.5265
18.1907 | -34.9789 | -28.5342
18.1968 | -34.9794 | -28.5417
18.2029 | -34.9799 | -28.5486
18.209 | -34.9804 | -28.5561
18.2151 | -34.9809 | -28.5636
18.2212 | -34.9814 | -28.5706
18.2273 | -34.9819 | -28.5778
18.2334 | -34.9824 | -28.5852
18.2395 | -34.9829 | -28.5926
18.2456 | -34.9833 | -28.5998
18.2517 | -34.9838 | -28.6071
18.2578 | -34.9843 | -28.6143
18.2639 | -34.9848 | -28.6218
18.2701 | -34.9853 | -28.6291
18.2762 | -34.9858 | -28.6366
18.2823 | -34.9863 | -28.6442
18.2884 | -34.9867 | -28.6518
18.2945 | -34.9872 | -28.659
18.3006 | -34.9877 | -28.6662
18.3067 | -34.9882 | -28.6735
18.3128 | -34.9887 | -28.6807
18.3189 | -34.9892 | -28.6883
18.325 | -34.9896 | -28.6953
18.3311 | -34.9901 | -28.7029
18.3372 | -34.9906 | -28.71
18.3433 | -34.991 | -28.7171
18.3494 | -34.9915 | -28.7244
18.3555 | -34.992 | -28.7313
18.3616 | -34.9924 | -28.7383
18.3677 | -34.9929 | -28.7452
18.3738 | -34.9934 | -28.7525
18.3799 | -34.9939 | -28.7596
18.386 | -34.9943 | -28.7668
18.3921 | -34.9948 | -28.7741
18.3982 | -34.9953 | -28.7814
18.4043 | -34.9958 | -28.7887
18.4105 | -34.9962 | -28.7959
18.4166 | -34.9967 | -28.8032
18.4227 | -34.9972 | -28.8105
18.4288 | -34.9977 | -28.818
18.4349 | -34.9982 | -28.8251
18.441 | -34.9987 | -28.8324
18.4471 | -34.9992 | -28.8399
18.4532 | -34.9997 | -28.8473
18.4593 | -35.0002 | -28.8545
18.4654 | -35.0007 | -28.8616
18.4715 | -35.0012 | -28.8688
18.4776 | -35.0017 | -28.8761
18.4837 | -35.0022 | -28.8835
18.4898 | -35.0027 | -28.891
18.4959 | -35.0032 | -28.8981
18.502 | -35.0038 | -28.9054
18.5081 | -35.0042 | -28.9129
18.5142 | -35.0047 | -28.9204
18.5203 | -35.0052 | -28.9273
18.5264 | -35.0056 | -28.9345
18.5325 | -35.0061 | -28.9418
18.5386 | -35.0067 | -28.9491
18.5447 | -35.0072 | -28.9564
18.5508 | -35.0077 | -28.9637
18.557 | -35.0081 | -28.971
18.5631 | -35.0086 | -28.9784
18.5692 | -35.0091 | -28.9857
18.5753 | -35.0096 | -28.9927
18.5814 | -35.0101 | -29.0002
18.5875 | -35.0106 | -29.0079
18.5936 | -35.0111 | -29.0155
18.5997 | -35.0116 | -29.0228
18.6058 | -35.0121 | -29.0299
18.6119 | -35.0126 | -29.0371
18.618 | -35.0131 | -29.0445
18.6241 | -35.0136 | -29.052
18.6302 | -35.0141 | -29.0594
18.6363 | -35.0146 | -29.0664
18.6424 | -35.0152 | -29.0736
18.6485 | -35.0156 | -29.081
18.6546 | -35.0161 | -29.0882
18.6607 | -35.0166 | -29.0954
18.6668 | -35.0171 | -29.1025
18.6729 | -35.0176 | -29.1098
18.679 | -35.0181 | -29.1168
18.6851 | -35.0186 | -29.124
18.6912 | -35.0191 | -29.1314
18.6974 | -35.0195 | -29.1388
18.7035 | -35.02 | -29.1459
18.7096 | -35.0205 | -29.153
18.7157 | -35.0211 | -29.1601
18.7218 | -35.0216 | -29.1674
18.7279 | -35.022 | -29.1744
18.734 | -35.0226 | -29.1815
18.7401 | -35.023 | -29.1885
18.7462 | -35.0235 | -29.1954
18.7523 | -35.0241 | -29.2024
18.7584 | -35.0245 | -29.2097
18.7645 | -35.0251 | -29.2168
18.7706 | -35.0255 | -29.2241
18.7767 | -35.0261 | -29.2313
18.7828 | -35.0266 | -29.2382
18.7889 | -35.027 | -29.2452
18.795 | -35.0275 | -29.2522
18.8011 | -35.028 | -29.2594
18.8072 | -35.0286 | -29.2666
18.8133 | -35.029 | -29.2739
18.8194 | -35.0296 | -29.2811
18.8255 | -35.0301 | -29.2883
18.8316 | -35.0305 | -29.2956
18.8377 | -35.031 | -29.3029
18.8439 | -35.0315 | -29.3099
18.85 | -35.032 | -29.3171
18.8561 | -35.0325 | -29.324
18.8622 | -35.033 | -29.3314
18.8683 | -35.0335 | -29.3385
18.8744 | -35.034 | -29.3457
18.8805 | -35.0345 | -29.3527
18.8866 | -35.035 | -29.3598
18.8927 | -35.0355 | -29.3665
18.8988 | -35.036 | -29.3737
18.9049 | -35.0365 | -29.381
18.911 | -35.037 | -29.3882
18.9171 | -35.0375 | -29.3954
18.9232 | -35.038 | -29.4023
18.9293 | -35.0385 | -29.4095
18.9354 | -35.039 | -29.4165
18.9415 | -35.0394 | -29.4235
18.9476 | -35.0399 | -29.4308
18.9537 | -35.0405 | -29.4381
18.9598 | -35.041 | -29.4448
18.9659 | -35.0415 | -29.4519
18.972 | -35.0419 | -29.4591
18.9781 | -35.0425 | -29.4662
18.9843 | -35.043 | -29.4733
18.9904 | -35.0434 | -29.4806
18.9965 | -35.0439 | -29.4877
19.0026 | -35.0444 | -29.4949
19.0087 | -35.0449 | -29.502
19.0148 | -35.0454 | -29.5088
19.0209 | -35.0459 | -29.5163
19.027 | -35.0464 | -29.5235
19.0331 | -35.047 | -29.5307
19.0392 | -35.0475 | -29.5378
19.0453 | -35.048 | -29.5449
19.0514 | -35.0484 | -29.5521
19.0575 | -35.0489 | -29.5591
19.0636 | -35.0494 | -29.566
19.0697 | -35.0499 | -29.5734
19.0758 | -35.0505 | -29.5804
19.0819 | -35.051 | -29.5878
19.088 | -35.0514 | -29.5947
19.0941 | -35.052 | -29.6015
19.1002 | -35.0525 | -29.6084
19.1063 | -35.0531 | -29.6158
19.1124 | -35.0536 | -29.6229
19.1185 | -35.0541 | -29.6304
19.1246 | -35.0545 | -29.6374
19.1308 | -35.0551 | -29.6441
19.1369 | -35.0556 | -29.6509
19.143 | -35.0561 | -29.658
19.1491 | -35.0566 | -29.6652
19.1552 | -35.0571 | -29.6726
19.1613 | -35.0576 | -29.6798
19.1674 | -35.0581 | -29.687
19.1735 | -35.0586 | -29.6939
19.1796 | -35.0591 | -29.701
19.1857 | -35.0596 | -29.7082
19.1918 | -35.0601 | -29.7151
19.1979 | -35.0606 | -29.722
19.204 | -35.0611 | -29.7289
19.2101 | -35.0616 | -29.7361
19.2162 | -35.0621 | -29.7434
19.2223 | -35.0626 | -29.7502
19.2284 | -35.0631 | -29.7576
19.2345 | -35.0637 | -29.7647
19.2406 | -35.0642 | -29.7718
19.2467 | -35.0647 | -29.779
19.2528 | -35.0652 | -29.786
19.2589 | -35.0657 | -29.7933
19.265 | -35.0662 | -29.8004
19.2712 | -35.0667 | -29.8076
19.2773 | -35.0672 | -29.8142
19.2834 | -35.0677 | -29.8212
19.2895 | -35.0682 | -29.8283
19.2956 | -35.0687 | -29.8356
19.3017 | -35.0692 | -29.8425
19.3078 | -35.0697 | -29.8496
19.3139 | -35.0703 | -29.8565
19.32 | -35.0708 | -29.8636
19.3261 | -35.0713 | -29.8707
19.3322 | -35.0718 | -29.8776
19.3383 | -35.0723 | -29.8846
19.3444 | -35.0728 | -29.8913
19.3505 | -35.0733 | -29.8984
19.3566 | -35.0738 | -29.9054
19.3627 | -35.0743 | -29.9127
19.3688 | -35.0748 | -29.9201
19.3749 | -35.0753 | -29.9271
19.381 | -35.0758 | -29.9338
19.3871 | -35.0762 | -29.941
19.3932 | -35.0768 | -29.9485
19.3993 | -35.0773 | -29.9555
19.4054 | -35.0778 | -29.9626
19.4115 | -35.0783 | -29.9696
19.4177 | -35.0789 | -29.9766
19.4238 | -35.0793 | -29.9835
19.4299 | -35.0799 | -29.9904
19.436 | -35.0804 | -29.9976
19.4421 | -35.0809 | -30.0044
19.4482 | -35.0814 | -30.0115
19.4543 | -35.0819 | -30.0184
19.4604 | -35.0824 | -30.0252
19.4665 | -35.0829 | -30.0324
19.4726 | -35.0834 | -30.0393
19.4787 | -35.0839 | -30.0463
19.4848 | -35.0844 | -30.0533
19.4909 | -35.085 | -30.06
19.497 | -35.0855 | -30.0671
19.5031 | -35.086 | -30.0742
19.5092 | -35.0865 | -30.0811
19.5153 | -35.087 | -30.0882
19.5214 | -35.0875 | -30.095
19.5275 | -35.088 | -30.1021
19.5336 | -35.0885 | -30.1093
19.5397 | -35.0891 | -30.1163
19.5458 | -35.0896 | -30.1232
19.5519 | -35.0901 | -30.1303
19.5581 | -35.0906 | -30.1371
19.5642 | -35.0912 | -30.1442
19.5703 | -35.0917 | -30.1511
19.5764 | -35.0922 | -30.158
19.5825 | -35.0927 | -30.1651
19.5886 | -35.0931 | -30.172
19.5947 | -35.0937 | -30.179
19.6008 | -35.0942 | -30.1863
19.6069 | -35.0947 | -30.1932
19.613 | -35.0952 | -30.2001
19.6191 | -35.0957 | -30.2072
19.6252 | -35.0962 | -30.2142
19.6313 | -35.0967 | -30.221
19.6374 | -35.0972 | -30.2281
19.6435 | -35.0977 | -30.2351
19.6496 | -35.0982 | -30.2421
19.6557 | -35.0988 | -30.249
19.6618 | -35.0993 | -30.256
19.6679 | -35.0998 | -30.263
19.674 | -35.1003 | -30.27
19.6801 | -35.1008 | -30.2768
19.6862 | -35.1013 | -30.2838
19.6923 | -35.1018 | -30.2906
19.6984 | -35.1023 | -30.2974
19.7046 | -35.1028 | -30.3044
19.7107 | -35.1033 | -30.3113
19.7168 | -35.1038 | -30.3184
19.7229 | -35.1043 | -30.3254
19.729 | -35.1048 | -30.3326
19.7351 | -35.1053 | -30.3393
19.7412 | -35.1058 | -30.3461
19.7473 | -35.1063 | -30.3531
19.7534 | -35.1068 | -30.3602
19.7595 | -35.1073 | -30.3673
19.7656 | -35.1079 | -30.3745
19.7717 | -35.1084 | -30.3813
19.7778 | -35.1089 | -30.3885
19.7839 | -35.1094 | -30.3954
19.79 | -35.1099 | -30.4022
19.7961 | -35.1104 | -30.4093
19.8022 | -35.1109 | -30.4164
19.8083 | -35.1114 | -30.4232
19.8144 | -35.1119 | -30.4304
19.8205 | -35.1123 | -30.4375
19.8266 | -35.1129 | -30.4447
19.8327 | -35.1133 | -30.4518
19.8388 | -35.1139 | -30.4588
19.845 | -35.1144 | -30.4657
19.8511 | -35.1149 | -30.4729
19.8572 | -35.1154 | -30.4802
19.8633 | -35.1159 | -30.4872
19.8694 | -35.1165 | -30.4944
19.8755 | -35.117 | -30.5015
19.8816 | -35.1176 | -30.5082
19.8877 | -35.1181 | -30.5152
19.8938 | -35.1186 | -30.5222
19.8999 | -35.1191 | -30.5289
19.906 | -35.1197 | -30.536
19.9121 | -35.1202 | -30.5425
19.9182 | -35.1207 | -30.5491
19.9243 | -35.1213 | -30.5563
19.9304 | -35.1217 | -30.5638
19.9365 | -35.1222 | -30.5708
19.9426 | -35.1227 | -30.5778
19.9487 | -35.1233 | -30.5844
19.9548 | -35.1238 | -30.5916
19.9609 | -35.1244 | -30.5985
19.967 | -35.1249 | -30.6053
19.9731 | -35.1253 | -30.6122
19.9792 | -35.1258 | -30.619
19.9853 | -35.1263 | -30.6261
19.9915 | -35.1268 | -30.6331
19.9976 | -35.1274 | -30.6399
20.0037 | -35.1279 | -30.6466
20.0098 | -35.1284 | -30.6537
20.0159 | -35.1289 | -30.6605
20.022 | -35.1294 | -30.6679
20.0281 | -35.1299 | -30.6752
20.0342 | -35.1304 | -30.6822
20.0403 | -35.1309 | -30.6888
20.0464 | -35.1314 | -30.6962
20.0525 | -35.1319 | -30.7032
20.0586 | -35.1325 | -30.7099
20.0647 | -35.133 | -30.7171
20.0708 | -35.1336 | -30.7241
20.0769 | -35.1341 | -30.7311
20.083 | -35.1346 | -30.738
20.0891 | -35.1351 | -30.7449
20.0952 | -35.1356 | -30.7517
20.1013 | -35.1361 | -30.7587
20.1074 | -35.1366 | -30.7658
20.1135 | -35.1371 | -30.7727
20.1196 | -35.1376 | -30.7796
20.1257 | -35.1382 | -30.7868
20.1319 | -35.1387 | -30.7937
20.138 | -35.1392 | -30.8005
20.1441 | -35.1397 | -30.8072
20.1502 | -35.1402 | -30.8144
20.1563 | -35.1407 | -30.8212
20.1624 | -35.1412 | -30.828
20.1685 | -35.1418 | -30.8351
20.1746 | -35.1423 | -30.8424
20.1807 | -35.1427 | -30.8493
20.1868 | -35.1432 | -30.8561
20.1929 | -35.1437 | -30.8629
20.199 | -35.1442 | -30.8699
20.2051 | -35.1448 | -30.8768
20.2112 | -35.1452 | -30.8832
20.2173 | -35.1457 | -30.8905
20.2234 | -35.1462 | -30.8977
20.2295 | -35.1467 | -30.9045
20.2356 | -35.1472 | -30.9114
20.2417 | -35.1478 | -30.9183
20.2478 | -35.1483 | -30.9252
20.2539 | -35.1487 | -30.9322
20.26 | -35.1493 | -30.9389
20.2661 | -35.1497 | -30.9459
20.2723 | -35.1502 | -30.9526
20.2784 | -35.1507 | -30.9597
20.2845 | -35.1512 | -30.9668
20.2906 | -35.1517 | -30.9738
20.2967 | -35.1523 | -30.9805
20.3028 | -35.1528 | -30.9873
20.3089 | -35.1533 | -30.9942
20.315 | -35.1538 | -31.0013
20.3211 | -35.1542 | -31.0084
20.3272 | -35.1548 | -31.0149
20.3333 | -35.1553 | -31.0214
20.3394 | -35.1558 | -31.0284
20.3455 | -35.1563 | -31.0352
20.3516 | -35.1568 | -31.0422
20.3577 | -35.1573 | -31.0492
20.3638 | -35.1578 | -31.0563
20.3699 | -35.1584 | -31.063
20.376 | -35.1589 | -31.0701
20.3821 | -35.1594 | -31.0767
20.3882 | -35.1599 | -31.0835
20.3943 | -35.1604 | -31.0905
20.4004 | -35.1609 | -31.0972
20.4065 | -35.1614 | -31.1041
20.4126 | -35.1619 | -31.111
20.4188 | -35.1625 | -31.118
20.4249 | -35.163 | -31.1251
20.431 | -35.1634 | -31.1324
20.4371 | -35.1639 | -31.1388
20.4432 | -35.1644 | -31.1459
20.4493 | -35.1649 | -31.1529
20.4554 | -35.1654 | -31.1597
20.4615 | -35.1659 | -31.1666
20.4676 | -35.1664 | -31.1736
20.4737 | -35.1669 | -31.1804
20.4798 | -35.1674 | -31.1876
20.4859 | -35.1679 | -31.1941
20.492 | -35.1684 | -31.2009
20.4981 | -35.1689 | -31.2077
20.5042 | -35.1694 | -31.2144
20.5103 | -35.1699 | -31.2213
20.5164 | -35.1704 | -31.2283
20.5225 | -35.1709 | -31.2355
20.5286 | -35.1714 | -31.2429
20.5347 | -35.1719 | -31.2496
20.5408 | -35.1724 | -31.2566
20.5469 | -35.173 | -31.2633
20.553 | -35.1735 | -31.2701
20.5592 | -35.1739 | -31.2768
20.5653 | -35.1744 | -31.2836
20.5714 | -35.1749 | -31.2902
20.5775 | -35.1754 | -31.2967
20.5836 | -35.1759 | -31.3038
20.5897 | -35.1764 | -31.3105
20.5958 | -35.1769 | -31.3173
20.6019 | -35.1774 | -31.324
20.608 | -35.1779 | -31.331
20.6141 | -35.1784 | -31.3374
20.6202 | -35.1789 | -31.3444
20.6263 | -35.1794 | -31.351
20.6324 | -35.1799 | -31.3576
20.6385 | -35.1804 | -31.3643
20.6446 | -35.1808 | -31.3714
20.6507 | -35.1814 | -31.3783
20.6568 | -35.1819 | -31.3851
20.6629 | -35.1824 | -31.3918
20.669 | -35.1829 | -31.3985
20.6751 | -35.1834 | -31.4053
20.6812 | -35.1839 | -31.4121
20.6873 | -35.1844 | -31.4189
20.6934 | -35.1849 | -31.4258
20.6995 | -35.1854 | -31.4325
20.7057 | -35.1859 | -31.4393
20.7118 | -35.1865 | -31.4461
20.7179 | -35.187 | -31.4528
20.724 | -35.1874 | -31.4594
20.7301 | -35.1879 | -31.4666
20.7362 | -35.1884 | -31.4736
20.7423 | -35.1889 | -31.4803
20.7484 | -35.1894 | -31.4871
20.7545 | -35.1899 | -31.4938
20.7606 | -35.1904 | -31.5005
20.7667 | -35.1909 | -31.5073
20.7728 | -35.1914 | -31.5141
20.7789 | -35.1919 | -31.5208
20.785 | -35.1924 | -31.5278
20.7911 | -35.1929 | -31.5341
20.7972 | -35.1933 | -31.5409
20.8033 | -35.1938 | -31.5471
20.8094 | -35.1942 | -31.5541
20.8155 | -35.1947 | -31.5609
20.8216 | -35.1952 | -31.568
20.8277 | -35.1957 | -31.5747
20.8338 | -35.1961 | -31.5814
20.8399 | -35.1966 | -31.5879
20.8461 | -35.1971 | -31.5947
20.8522 | -35.1976 | -31.6011
20.8583 | -35.1981 | -31.6077
20.8644 | -35.1986 | -31.6143
20.8705 | -35.1991 | -31.6212
20.8766 | -35.1996 | -31.6282
20.8827 | -35.2001 | -31.635
20.8888 | -35.2006 | -31.6417
20.8949 | -35.2011 | -31.6484
20.901 | -35.2016 | -31.655
20.9071 | -35.2021 | -31.662
20.9132 | -35.2026 | -31.6687
20.9193 | -35.2031 | -31.6754
20.9254 | -35.2036 | -31.6822
20.9315 | -35.2041 | -31.6887
20.9376 | -35.2045 | -31.6959
20.9437 | -35.2051 | -31.7029
20.9498 | -35.2055 | -31.7095
20.9559 | -35.2061 | -31.716
20.962 | -35.2066 | -31.7231
20.9681 | -35.2071 | -31.7296
20.9742 | -35.2076 | -31.7361
20.9803 | -35.2081 | -31.7428
20.9864 | -35.2086 | -31.7495
20.9926 | -35.2091 | -31.756
20.9987 | -35.2096 | -31.7625
21.0048 | -35.2101 | -31.7693
21.0109 | -35.2105 | -31.776
21.017 | -35.211 | -31.7828
21.0231 | -35.2114 | -31.7896
21.0292 | -35.2119 | -31.7964
21.0353 | -35.2124 | -31.8028
21.0414 | -35.2129 | -31.8096
21.0475 | -35.2134 | -31.8165
21.0536 | -35.2139 | -31.8232
21.0597 | -35.2144 | -31.8298
21.0658 | -35.2148 | -31.8365
21.0719 | -35.2152 | -31.8432
21.078 | -35.2158 | -31.8503
21.0841 | -35.2163 | -31.8572
21.0902 | -35.2168 | -31.8639
21.0963 | -35.2173 | -31.8705
21.1024 | -35.2177 | -31.8771
21.1085 | -35.2182 | -31.8831
21.1146 | -35.2187 | -31.89
21.1207 | -35.2191 | -31.8966
21.1268 | -35.2197 | -31.9033
21.133 | -35.2202 | -31.9098
21.1391 | -35.2206 | -31.9165
21.1452 | -35.2211 | -31.9237
21.1513 | -35.2216 | -31.9305
21.1574 | -35.2221 | -31.9375
21.1635 | -35.2226 | -31.9438
21.1696 | -35.223 | -31.9504
21.1757 | -35.2235 | -31.9573
21.1818 | -35.224 | -31.964
21.1879 | -35.2245 | -31.9707
21.194 | -35.225 | -31.9773
21.2001 | -35.2255 | -31.9839
21.2062 | -35.226 | -31.9907
21.2123 | -35.2264 | -31.9973
21.2184 | -35.2269 | -32.0042
21.2245 | -35.2274 | -32.011
21.2306 | -35.2278 | -32.0176
21.2367 | -35.2283 | -32.0243
21.2428 | -35.2287 | -32.031
21.2489 | -35.2292 | -32.0376
21.255 | -35.2297 | -32.0444
21.2611 | -35.2301 | -32.0512
21.2672 | -35.2306 | -32.0579
21.2733 | -35.2311 | -32.0645
21.2795 | -35.2316 | -32.0712
21.2856 | -35.2321 | -32.0778
21.2917 | -35.2325 | -32.0847
21.2978 | -35.233 | -32.0916
21.3039 | -35.2334 | -32.0982
21.31 | -35.2339 | -32.1045
21.3161 | -35.2344 | -32.1111
21.3222 | -35.2349 | -32.1173
21.3283 | -35.2354 | -32.1241
21.3344 | -35.2358 | -32.131
21.3405 | -35.2363 | -32.1379
21.3466 | -35.2368 | -32.1446
21.3527 | -35.2372 | -32.1508
21.3588 | -35.2377 | -32.1572
21.3649 | -35.2381 | -32.1637
21.371 | -35.2386 | -32.1703
21.3771 | -35.2391 | -32.1767
21.3832 | -35.2395 | -32.1833
21.3893 | -35.24 | -32.1895
21.3954 | -35.2405 | -32.1959
21.4015 | -35.2409 | -32.2027
21.4076 | -35.2414 | -32.2093
21.4137 | -35.2419 | -32.2158
21.4199 | -35.2423 | -32.2224
21.426 | -35.2428 | -32.2292
21.4321 | -35.2433 | -32.2358
21.4382 | -35.2437 | -32.2423
21.4443 | -35.2442 | -32.249
21.4504 | -35.2447 | -32.2557
21.4565 | -35.2452 | -32.2621
21.4626 | -35.2457 | -32.2687
21.4687 | -35.2462 | -32.2751
21.4748 | -35.2467 | -32.2817
21.4809 | -35.2471 | -32.2884
21.487 | -35.2476 | -32.2951
21.4931 | -35.248 | -32.3016
21.4992 | -35.2485 | -32.3078
21.5053 | -35.2489 | -32.3145
21.5114 | -35.2494 | -32.321
21.5175 | -35.2499 | -32.3275
21.5236 | -35.2503 | -32.3343
21.5297 | -35.2508 | -32.3408
21.5358 | -35.2513 | -32.3473
21.5419 | -35.2517 | -32.354
21.548 | -35.2522 | -32.3607
21.5541 | -35.2526 | -32.3675
21.5602 | -35.2531 | -32.3741
21.5664 | -35.2536 | -32.3804
21.5725 | -35.254 | -32.3868
21.5786 | -35.2544 | -32.3931
21.5847 | -35.2549 | -32.3995
21.5908 | -35.2553 | -32.4061
21.5969 | -35.2557 | -32.4125
21.603 | -35.2562 | -32.4188
21.6091 | -35.2567 | -32.4255
21.6152 | -35.2572 | -32.4321
21.6213 | -35.2577 | -32.4389
21.6274 | -35.2582 | -32.4454
21.6335 | -35.2587 | -32.452
21.6396 | -35.2591 | -32.4584
21.6457 | -35.2596 | -32.4651
21.6518 | -35.26 | -32.4716
21.6579 | -35.2605 | -32.4777
21.664 | -35.261 | -32.4842
21.6701 | -35.2614 | -32.4905
21.6762 | -35.2618 | -32.4971
21.6823 | -35.2623 | -32.5037
21.6884 | -35.2628 | -32.5102
21.6945 | -35.2632 | -32.5167
21.7006 | -35.2637 | -32.5232
21.7068 | -35.2642 | -32.5295
21.7129 | -35.2646 | -32.5362
21.719 | -35.2651 | -32.5429
21.7251 | -35.2656 | -32.5495
21.7312 | -35.266 | -32.5559
21.7373 | -35.2665 | -32.5623
21.7434 | -35.267 | -32.5689
21.7495 | -35.2674 | -32.5753
21.7556 | -35.2678 | -32.582
21.7617 | -35.2683 | -32.5882
21.7678 | -35.2688 | -32.5949
21.7739 | -35.2692 | -32.601
21.78 | -35.2697 | -32.6076
21.7861 | -35.2702 | -32.6141
21.7922 | -35.2706 | -32.6206
21.7983 | -35.2711 | -32.6273
21.8044 | -35.2716 | -32.634
21.8105 | -35.2721 | -32.6403
21.8166 | -35.2725 | -32.6467
21.8227 | -35.273 | -32.6531
21.8288 | -35.2734 | -32.6594
21.8349 | -35.2739 | -32.6656
21.841 | -35.2744 | -32.6723
21.8471 | -35.2748 | -32.6788
21.8533 | -35.2752 | -32.6849
21.8594 | -35.2757 | -32.6918
21.8655 | -35.2762 | -32.6985
21.8716 | -35.2766 | -32.7051
21.8777 | -35.2771 | -32.7113
21.8838 | -35.2775 | -32.7176
21.8899 | -35.278 | -32.724
21.896 | -35.2784 | -32.7306
21.9021 | -35.2789 | -32.737
21.9082 | -35.2793 | -32.7434
21.9143 | -35.2797 | -32.7497
21.9204 | -35.2802 | -32.7563
21.9265 | -35.2806 | -32.7625
21.9326 | -35.2811 | -32.7693
21.9387 | -35.2815 | -32.7757
21.9448 | -35.282 | -32.7823
21.9509 | -35.2824 | -32.7886
21.957 | -35.2829 | -32.795
21.9631 | -35.2834 | -32.8018
21.9692 | -35.2838 | -32.8084
21.9753 | -35.2843 | -32.8149
21.9814 | -35.2847 | -32.8212
21.9875 | -35.2851 | -32.8277
21.9937 | -35.2856 | -32.8343
21.9998 | -35.2861 | -32.8408
22.0059 | -35.2866 | -32.8475
22.012 | -35.287 | -32.8539
22.0181 | -35.2875 | -32.8601
22.0242 | -35.2879 | -32.8666
22.0303 | -35.2884 | -32.8731
22.0364 | -35.2888 | -32.8796
22.0425 | -35.2893 | -32.8859
22.0486 | -35.2897 | -32.8925
22.0547 | -35.2901 | -32.8988
22.0608 | -35.2905 | -32.9053
22.0669 | -35.291 | -32.9116
22.073 | -35.2915 | -32.9179
22.0791 | -35.2919 | -32.9242
22.0852 | -35.2924 | -32.9306
22.0913 | -35.2928 | -32.937
22.0974 | -35.2933 | -32.9432
22.1035 | -35.2937 | -32.9496
22.1096 | -35.2942 | -32.9562
22.1157 | -35.2946 | -32.9625
22.1218 | -35.2951 | -32.9691
22.1279 | -35.2954 | -32.9754
22.134 | -35.2959 | -32.9817
22.1402 | -35.2963 | -32.9882
22.1463 | -35.2968 | -32.9946
22.1524 | -35.2972 | -33.001
22.1585 | -35.2977 | -33.0074
22.1646 | -35.2981 | -33.0136
22.1707 | -35.2986 | -33.0202
22.1768 | -35.2991 | -33.0267
22.1829 | -35.2995 | -33.033
22.189 | -35.2999 | -33.0392
22.1951 | -35.3004 | -33.0455
22.2012 | -35.3008 | -33.0519
22.2073 | -35.3012 | -33.0583
22.2134 | -35.3017 | -33.0647
22.2195 | -35.3021 | -33.0712
22.2256 | -35.3026 | -33.0775
22.2317 | -35.303 | -33.0835
22.2378 | -35.3034 | -33.0899
22.2439 | -35.3039 | -33.0962
22.25 | -35.3044 | -33.1027
22.2561 | -35.3048 | -33.1093
22.2622 | -35.3052 | -33.1158
22.2683 | -35.3057 | -33.1218
22.2744 | -35.3061 | -33.1281
22.2806 | -35.3065 | -33.1342
22.2867 | -35.307 | -33.1407
22.2928 | -35.3074 | -33.147
22.2989 | -35.3079 | -33.1534
22.305 | -35.3083 | -33.1592
22.3111 | -35.3087 | -33.1655
22.3172 | -35.3092 | -33.1718
22.3233 | -35.3096 | -33.178
22.3294 | -35.3101 | -33.1844
22.3355 | -35.3105 | -33.1903
22.3416 | -35.311 | -33.1966
22.3477 | -35.3115 | -33.2031
22.3538 | -35.3119 | -33.2093
22.3599 | -35.3124 | -33.2156
22.366 | -35.3128 | -33.2221
22.3721 | -35.3132 | -33.2282
22.3782 | -35.3137 | -33.2347
22.3843 | -35.3141 | -33.2411
22.3904 | -35.3146 | -33.2477
22.3965 | -35.315 | -33.2541
22.4026 | -35.3155 | -33.2602
22.4087 | -35.3158 | -33.2665
22.4148 | -35.3163 | -33.2727
22.4209 | -35.3167 | -33.279
22.4271 | -35.3172 | -33.2854
22.4332 | -35.3176 | -33.2915
22.4393 | -35.318 | -33.2977
22.4454 | -35.3184 | -33.3035
22.4515 | -35.3188 | -33.3098
22.4576 | -35.3193 | -33.3161
22.4637 | -35.3197 | -33.3226
22.4698 | -35.3201 | -33.3287
22.4759 | -35.3205 | -33.3347
22.482 | -35.321 | -33.3408
22.4881 | -35.3214 | -33.3472
22.4942 | -35.3219 | -33.3533
22.5003 | -35.3223 | -33.3596
22.5064 | -35.3227 | -33.3657
22.5125 | -35.3231 | -33.3721
22.5186 | -35.3235 | -33.3784
22.5247 | -35.324 | -33.3848
22.5308 | -35.3244 | -33.3911
22.5369 | -35.3249 | -33.3977
22.543 | -35.3253 | -33.4042
22.5491 | -35.3257 | -33.4101
22.5552 | -35.3261 | -33.4163
22.5613 | -35.3265 | -33.4226
22.5675 | -35.3269 | -33.429
22.5736 | -35.3273 | -33.4355
22.5797 | -35.3278 | -33.4416
22.5858 | -35.3282 | -33.4478
22.5919 | -35.3286 | -33.454
22.598 | -35.329 | -33.46
22.6041 | -35.3294 | -33.466
22.6102 | -35.3299 | -33.4722
22.6163 | -35.3303 | -33.4781
22.6224 | -35.3307 | -33.4845
22.6285 | -35.3311 | -33.4907
22.6346 | -35.3316 | -33.4966
22.6407 | -35.332 | -33.5031
22.6468 | -35.3324 | -33.5093
22.6529 | -35.3328 | -33.5152
22.659 | -35.3332 | -33.5215
22.6651 | -35.3336 | -33.5278
22.6712 | -35.334 | -33.5341
22.6773 | -35.3345 | -33.5403
22.6834 | -35.3349 | -33.5465
22.6895 | -35.3353 | -33.5528
22.6956 | -35.3357 | -33.5593
22.7017 | -35.3361 | -33.5655
22.7079 | -35.3365 | -33.5722
22.714 | -35.3369 | -33.5784
22.7201 | -35.3374 | -33.5843
22.7262 | -35.3378 | -33.5905
22.7323 | -35.3382 | -33.5967
22.7384 | -35.3387 | -33.6031
22.7445 | -35.3391 | -33.6092
22.7506 | -35.3395 | -33.6156
22.7567 | -35.3399 | -33.6218
22.7628 | -35.3403 | -33.6277
22.7689 | -35.3407 | -33.634
22.775 | -35.3411 | -33.6403
22.7811 | -35.3416 | -33.6464
22.7872 | -35.342 | -33.6527
22.7933 | -35.3424 | -33.6587
22.7994 | -35.3428 | -33.6649
22.8055 | -35.3432 | -33.6711
22.8116 | -35.3436 | -33.6774
22.8177 | -35.344 | -33.6837
22.8238 | -35.3445 | -33.6901
22.8299 | -35.3448 | -33.6961
22.836 | -35.3452 | -33.7025
22.8421 | -35.3457 | -33.7086
22.8482 | -35.3462 | -33.7148
22.8544 | -35.3466 | -33.7211
22.8605 | -35.347 | -33.7274
22.8666 | -35.3474 | -33.7334
22.8727 | -35.3478 | -33.7398
22.8788 | -35.3482 | -33.7459
22.8849 | -35.3486 | -33.752
22.891 | -35.349 | -33.7583
22.8971 | -35.3494 | -33.7644
22.9032 | -35.3498 | -33.7705
22.9093 | -35.3502 | -33.7767
22.9154 | -35.3506 | -33.7829
22.9215 | -35.351 | -33.7891
22.9276 | -35.3515 | -33.7954
22.9337 | -35.3518 | -33.8013
22.9398 | -35.3523 | -33.8075
22.9459 | -35.3527 | -33.8136
22.952 | -35.3531 | -33.82
22.9581 | -35.3535 | -33.8262
22.9642 | -35.3539 | -33.8325
22.9703 | -35.3544 | -33.8384
22.9764 | -35.3548 | -33.8445
22.9825 | -35.3552 | -33.8507
22.9886 | -35.3556 | -33.8567
22.9948 | -35.356 | -33.8629
23.0009 | -35.3564 | -33.8689
23.007 | -35.3568 | -33.875
23.0131 | -35.3573 | -33.8811
23.0192 | -35.3577 | -33.8873
23.0253 | -35.3581 | -33.8934
23.0314 | -35.3586 | -33.8994
23.0375 | -35.359 | -33.9057
23.0436 | -35.3594 | -33.9116
23.0497 | -35.3598 | -33.9175
23.0558 | -35.3603 | -33.9238
23.0619 | -35.3607 | -33.93
23.068 | -35.3612 | -33.9361
23.0741 | -35.3616 | -33.9421
23.0802 | -35.362 | -33.9481
23.0863 | -35.3624 | -33.9545
23.0924 | -35.3628 | -33.9607
23.0985 | -35.3632 | -33.9669
23.1046 | -35.3636 | -33.9732
23.1107 | -35.364 | -33.9793
23.1168 | -35.3644 | -33.9855
23.1229 | -35.3649 | -33.9917
23.129 | -35.3652 | -33.9981
23.1351 | -35.3657 | -34.0043
23.1413 | -35.3661 | -34.0105
23.1474 | -35.3665 | -34.0164
23.1535 | -35.3669 | -34.0225
23.1596 | -35.3673 | -34.0285
23.1657 | -35.3677 | -34.0344
23.1718 | -35.3681 | -34.0406
23.1779 | -35.3685 | -34.0467
23.184 | -35.369 | -34.0525
23.1901 | -35.3694 | -34.0586
23.1962 | -35.3697 | -34.0648
23.2023 | -35.3702 | -34.0707
23.2084 | -35.3706 | -34.0769
23.2145 | -35.371 | -34.0831
23.2206 | -35.3714 | -34.0892
23.2267 | -35.3718 | -34.0953
23.2328 | -35.3722 | -34.1014
23.2389 | -35.3727 | -34.1077
23.245 | -35.3731 | -34.1137
23.2511 | -35.3735 | -34.1196
23.2572 | -35.3739 | -34.1256
23.2633 | -35.3743 | -34.1316
23.2694 | -35.3747 | -34.138
23.2755 | -35.3751 | -34.1442
23.2817 | -35.3755 | -34.1504
23.2878 | -35.376 | -34.1564
23.2939 | -35.3764 | -34.1625
23.3 | -35.3768 | -34.1687
23.3061 | -35.3772 | -34.1746
23.3122 | -35.3776 | -34.1806
23.3183 | -35.3779 | -34.1866
23.3244 | -35.3784 | -34.1925
23.3305 | -35.3788 | -34.1986
23.3366 | -35.3792 | -34.2046
23.3427 | -35.3795 | -34.2106
23.3488 | -35.3799 | -34.2168
23.3549 | -35.3804 | -34.2228
23.361 | -35.3808 | -34.2287
23.3671 | -35.3812 | -34.2349
23.3732 | -35.3816 | -34.2408
23.3793 | -35.382 | -34.2469
23.3854 | -35.3824 | -34.2531
23.3915 | -35.3828 | -34.2588
23.3976 | -35.3833 | -34.2647
23.4037 | -35.3836 | -34.2707
23.4098 | -35.384 | -34.277
23.4159 | -35.3845 | -34.2828
23.422 | -35.3848 | -34.2887
23.4282 | -35.3852 | -34.2945
23.4343 | -35.3856 | -34.3004
23.4404 | -35.386 | -34.3065
23.4465 | -35.3864 | -34.3127
23.4526 | -35.3868 | -34.3186
23.4587 | -35.3872 | -34.3246
23.4648 | -35.3876 | -34.3305
23.4709 | -35.388 | -34.3367
23.477 | -35.3884 | -34.3431
23.4831 | -35.3888 | -34.349
23.4892 | -35.3892 | -34.3553
23.4953 | -35.3896 | -34.3614
23.5014 | -35.39 | -34.3671
23.5075 | -35.3904 | -34.3731
23.5136 | -35.3908 | -34.379
23.5197 | -35.3912 | -34.3851
23.5258 | -35.3916 | -34.3912
23.5319 | -35.392 | -34.397
23.538 | -35.3924 | -34.403
23.5441 | -35.3928 | -34.4091
23.5502 | -35.3932 | -34.4154
23.5563 | -35.3936 | -34.4213
23.5624 | -35.394 | -34.4275
23.5686 | -35.3944 | -34.4337
23.5747 | -35.3948 | -34.4397
23.5808 | -35.3952 | -34.4458
23.5869 | -35.3956 | -34.452
23.593 | -35.396 | -34.4581
23.5991 | -35.3964 | -34.464
23.6052 | -35.3969 | -34.4699
23.6113 | -35.3973 | -34.4757
23.6174 | -35.3977 | -34.4817
23.6235 | -35.3981 | -34.4877
23.6296 | -35.3985 | -34.494
23.6357 | -35.3989 | -34.5003
23.6418 | -35.3993 | -34.5059
23.6479 | -35.3997 | -34.5122
23.654 | -35.4001 | -34.5184
23.6601 | -35.4005 | -34.5244
23.6662 | -35.4009 | -34.5306
23.6723 | -35.4013 | -34.5365
23.6784 | -35.4017 | -34.5426
23.6845 | -35.4021 | -34.5488
23.6906 | -35.4025 | -34.555
23.6967 | -35.4029 | -34.5611
23.7028 | -35.4033 | -34.5673
23.7089 | -35.4037 | -34.5736
23.7151 | -35.4041 | -34.5796
23.7212 | -35.4045 | -34.5855
23.7273 | -35.4049 | -34.5916
23.7334 | -35.4053 | -34.5975
23.7395 | -35.4057 | -34.6032
23.7456 | -35.4061 | -34.6092
23.7517 | -35.4065 | -34.6152
23.7578 | -35.4068 | -34.6213
23.7639 | -35.4072 | -34.6273
23.77 | -35.4076 | -34.6332
23.7761 | -35.4081 | -34.6389
23.7822 | -35.4085 | -34.6451
23.7883 | -35.4089 | -34.6511
23.7944 | -35.4093 | -34.6574
23.8005 | -35.4097 | -34.6635
23.8066 | -35.4101 | -34.6694
23.8127 | -35.4105 | -34.6755
23.8188 | -35.411 | -34.6816
23.8249 | -35.4113 | -34.6878
23.831 | -35.4117 | -34.6939
23.8371 | -35.4121 | -34.6997
23.8432 | -35.4125 | -34.7057
23.8493 | -35.4129 | -34.7111
23.8555 | -35.4133 | -34.7172
23.8616 | -35.4137 | -34.7234
23.8677 | -35.4141 | -34.7292
23.8738 | -35.4145 | -34.7355
23.8799 | -35.4149 | -34.7414
23.886 | -35.4153 | -34.7472
23.8921 | -35.4157 | -34.7534
23.8982 | -35.4161 | -34.7594
23.9043 | -35.4165 | -34.7656
23.9104 | -35.4169 | -34.7716
23.9165 | -35.4173 | -34.7775
23.9226 | -35.4176 | -34.7833
23.9287 | -35.418 | -34.7893
23.9348 | -35.4185 | -34.7953
23.9409 | -35.4189 | -34.8013
23.947 | -35.4192 | -34.8073
23.9531 | -35.4196 | -34.8129
23.9592 | -35.42 | -34.8189
23.9653 | -35.4204 | -34.825
23.9714 | -35.4207 | -34.8311
23.9775 | -35.4211 | -34.8372
23.9836 | -35.4216 | -34.8431
23.9897 | -35.422 | -34.8487
23.9958 | -35.4224 | -34.8547
24.002 | -35.4228 | -34.8608
24.0081 | -35.4232 | -34.8668
24.0142 | -35.4236 | -34.8729
24.0203 | -35.424 | -34.8788
24.0264 | -35.4244 | -34.8845
24.0325 | -35.4248 | -34.8905
24.0386 | -35.4251 | -34.8964
24.0447 | -35.4255 | -34.9023
24.0508 | -35.4259 | -34.9083
24.0569 | -35.4263 | -34.9142
24.063 | -35.4267 | -34.9199
24.0691 | -35.4271 | -34.9257
24.0752 | -35.4275 | -34.9315
24.0813 | -35.4279 | -34.9377
24.0874 | -35.4283 | -34.9435
24.0935 | -35.4287 | -34.9496
24.0996 | -35.4291 | -34.9551
24.1057 | -35.4295 | -34.961
24.1118 | -35.4298 | -34.967
24.1179 | -35.4302 | -34.9728
24.124 | -35.4306 | -34.9787
24.1301 | -35.431 | -34.9846
24.1362 | -35.4314 | -34.9907
24.1424 | -35.4318 | -34.9965
24.1485 | -35.4322 | -35.0022
24.1546 | -35.4326 | -35.0082
24.1607 | -35.433 | -35.0139
24.1668 | -35.4334 | -35.02
24.1729 | -35.4338 | -35.0258
24.179 | -35.4341 | -35.0317
24.1851 | -35.4346 | -35.0379
24.1912 | -35.4349 | -35.0441
24.1973 | -35.4354 | -35.0502
24.2034 | -35.4358 | -35.0558
24.2095 | -35.4362 | -35.0617
24.2156 | -35.4366 | -35.0678
24.2217 | -35.437 | -35.0736
24.2278 | -35.4374 | -35.0796
24.2339 | -35.4378 | -35.0853
24.24 | -35.4382 | -35.0913
24.2461 | -35.4386 | -35.0972
24.2522 | -35.439 | -35.1031
24.2583 | -35.4394 | -35.1093
24.2644 | -35.4398 | -35.1151
24.2705 | -35.4402 | -35.121
24.2766 | -35.4406 | -35.1268
24.2827 | -35.4409 | -35.1327
24.2889 | -35.4413 | -35.1384
24.295 | -35.4417 | -35.1445
24.3011 | -35.4421 | -35.1505
24.3072 | -35.4425 | -35.1562
24.3133 | -35.4429 | -35.1618
24.3194 | -35.4433 | -35.1679
24.3255 | -35.4438 | -35.1737
24.3316 | -35.4441 | -35.1798
24.3377 | -35.4445 | -35.1857
24.3438 | -35.4448 | -35.1919
24.3499 | -35.4452 | -35.198
24.356 | -35.4456 | -35.2042
24.3621 | -35.4459 | -35.2102
24.3682 | -35.4463 | -35.2159
24.3743 | -35.4467 | -35.2218
24.3804 | -35.4471 | -35.2277
24.3865 | -35.4474 | -35.2337
24.3926 | -35.4478 | -35.2395
24.3987 | -35.4482 | -35.2451
24.4048 | -35.4486 | -35.2511
24.4109 | -35.449 | -35.2569
24.417 | -35.4494 | -35.2628
24.4231 | -35.4497 | -35.2688
24.4293 | -35.4501 | -35.2745
24.4354 | -35.4505 | -35.2803
24.4415 | -35.4509 | -35.2861
24.4476 | -35.4512 | -35.2917
24.4537 | -35.4516 | -35.2976
24.4598 | -35.452 | -35.3036
24.4659 | -35.4524 | -35.3093
24.472 | -35.4528 | -35.3154
24.4781 | -35.4532 | -35.3215
24.4842 | -35.4535 | -35.3272
24.4903 | -35.4539 | -35.333
24.4964 | -35.4543 | -35.3392
24.5025 | -35.4547 | -35.3452
24.5086 | -35.4551 | -35.351
24.5147 | -35.4555 | -35.3566
24.5208 | -35.4559 | -35.3622
24.5269 | -35.4563 | -35.3681
24.533 | -35.4567 | -35.3738
24.5391 | -35.4571 | -35.3799
24.5452 | -35.4575 | -35.3858
24.5513 | -35.4579 | -35.3917
24.5574 | -35.4583 | -35.3976
24.5635 | -35.4587 | -35.4033
24.5696 | -35.4591 | -35.4092
24.5758 | -35.4595 | -35.4149
24.5819 | -35.4599 | -35.4209
24.588 | -35.4602 | -35.4269
24.5941 | -35.4606 | -35.4327
24.6002 | -35.461 | -35.4386
24.6063 | -35.4614 | -35.4443
24.6124 | -35.4618 | -35.4505
24.6185 | -35.4621 | -35.4564
24.6246 | -35.4625 | -35.4621
24.6307 | -35.4629 | -35.4682
24.6368 | -35.4633 | -35.4738
24.6429 | -35.4637 | -35.4795
24.649 | -35.4641 | -35.485
24.6551 | -35.4645 | -35.491
24.6612 | -35.4648 | -35.4965
24.6673 | -35.4652 | -35.5023
24.6734 | -35.4656 | -35.5081
24.6795 | -35.466 | -35.5141
24.6856 | -35.4664 | -35.5201
24.6917 | -35.4668 | -35.5259
24.6978 | -35.4672 | -35.5315
24.7039 | -35.4676 | -35.5374
24.71 | -35.468 | -35.5432
24.7162 | -35.4683 | -35.549
24.7223 | -35.4687 | -35.5547
24.7284 | -35.4691 | -35.5603
24.7345 | -35.4695 | -35.5663
24.7406 | -35.4699 | -35.5721
24.7467 | -35.4703 | -35.578
24.7528 | -35.4706 | -35.5839
24.7589 | -35.471 | -35.5895
24.765 | -35.4713 | -35.595
24.7711 | -35.4717 | -35.6007
24.7772 | -35.4721 | -35.6067
24.7833 | -35.4725 | -35.6127
24.7894 | -35.4729 | -35.6186
24.7955 | -35.4732 | -35.6245
24.8016 | -35.4737 | -35.6303
24.8077 | -35.4741 | -35.6361
24.8138 | -35.4745 | -35.6418
24.8199 | -35.4748 | -35.6477
24.826 | -35.4752 | -35.6534
24.8321 | -35.4755 | -35.6593
24.8382 | -35.4759 | -35.665
24.8443 | -35.4763 | -35.6709
24.8504 | -35.4766 | -35.6764
24.8565 | -35.477 | -35.6821
24.8627 | -35.4774 | -35.6881
24.8688 | -35.4778 | -35.6936
24.8749 | -35.4782 | -35.6991
24.881 | -35.4786 | -35.705
24.8871 | -35.479 | -35.711
24.8932 | -35.4794 | -35.7165
24.8993 | -35.4798 | -35.7222
24.9054 | -35.4802 | -35.7279
24.9115 | -35.4806 | -35.7338
24.9176 | -35.481 | -35.7398
24.9237 | -35.4814 | -35.7456
24.9298 | -35.4818 | -35.7515
24.9359 | -35.4822 | -35.7573
24.942 | -35.4826 | -35.763
24.9481 | -35.483 | -35.7689
24.9542 | -35.4834 | -35.7749
24.9603 | -35.4837 | -35.7807
24.9664 | -35.4841 | -35.7865
24.9725 | -35.4845 | -35.7925
24.9786 | -35.4848 | -35.7984
24.9847 | -35.4852 | -35.8042
24.9908 | -35.4856 | -35.8101
24.9969 | -35.486 | -35.816
25.0031 | -35.4863 | -35.8219
25.0092 | -35.4867 | -35.8278
25.0153 | -35.4871 | -35.8334
25.0214 | -35.4875 | -35.8393
25.0275 | -35.4879 | -35.8451
25.0336 | -35.4883 | -35.8509
25.0397 | -35.4887 | -35.8568
25.0458 | -35.489 | -35.8625
25.0519 | -35.4894 | -35.8682
25.058 | -35.4898 | -35.8741
25.0641 | -35.4902 | -35.8798
25.0702 | -35.4906 | -35.8858
25.0763 | -35.491 | -35.8916
25.0824 | -35.4914 | -35.8975
25.0885 | -35.4917 | -35.9033
25.0946 | -35.4921 | -35.909
25.1007 | -35.4925 | -35.9149
25.1068 | -35.4928 | -35.9205
25.1129 | -35.4932 | -35.9264
25.119 | -35.4936 | -35.9324
25.1251 | -35.494 | -35.9382
25.1312 | -35.4943 | -35.9439
25.1373 | -35.4947 | -35.9499
25.1435 | -35.4951 | -35.9557
25.1496 | -35.4955 | -35.9613
25.1557 | -35.4958 | -35.9671
25.1618 | -35.4962 | -35.9731
25.1679 | -35.4966 | -35.9789
25.174 | -35.497 | -35.9845
25.1801 | -35.4974 | -35.9904
25.1862 | -35.4978 | -35.9961
25.1923 | -35.4981 | -36.002
25.1984 | -35.4985 | -36.0077
25.2045 | -35.4989 | -36.0137
25.2106 | -35.4993 | -36.0194
25.2167 | -35.4997 | -36.0252
25.2228 | -35.5001 | -36.0308
25.2289 | -35.5004 | -36.0366
25.235 | -35.5008 | -36.0423
25.2411 | -35.5011 | -36.048
25.2472 | -35.5015 | -36.0539
25.2533 | -35.5019 | -36.0597
25.2594 | -35.5023 | -36.0655
25.2655 | -35.5026 | -36.0714
25.2716 | -35.503 | -36.0774
25.2777 | -35.5034 | -36.0834
25.2838 | -35.5037 | -36.0892
25.29 | -35.5041 | -36.095
25.2961 | -35.5045 | -36.101
25.3022 | -35.5048 | -36.1067
25.3083 | -35.5052 | -36.1124
25.3144 | -35.5056 | -36.1182
25.3205 | -35.5059 | -36.124
25.3266 | -35.5063 | -36.1297
25.3327 | -35.5067 | -36.1354
25.3388 | -35.5071 | -36.1413
25.3449 | -35.5075 | -36.1473
25.351 | -35.5078 | -36.1531
25.3571 | -35.5082 | -36.159
25.3632 | -35.5086 | -36.1646
25.3693 | -35.5089 | -36.1703
25.3754 | -35.5093 | -36.1762
25.3815 | -35.5097 | -36.182
25.3876 | -35.5101 | -36.1878
25.3937 | -35.5105 | -36.1935
25.3998 | -35.5108 | -36.1993
25.4059 | -35.5112 | -36.2053
25.412 | -35.5116 | -36.211
25.4181 | -35.5119 | -36.217
25.4242 | -35.5123 | -36.2229
25.4304 | -35.5126 | -36.2286
25.4365 | -35.513 | -36.2345
25.4426 | -35.5134 | -36.2404
25.4487 | -35.5138 | -36.2463
25.4548 | -35.5142 | -36.2524
25.4609 | -35.5146 | -36.2581
25.467 | -35.5149 | -36.2637
25.4731 | -35.5153 | -36.2696
25.4792 | -35.5157 | -36.2753
25.4853 | -35.5161 | -36.2808
25.4914 | -35.5164 | -36.2866
25.4975 | -35.5168 | -36.2924
25.5036 | -35.5172 | -36.2979
25.5097 | -35.5175 | -36.3035
25.5158 | -35.5179 | -36.3093
25.5219 | -35.5183 | -36.3149
25.528 | -35.5187 | -36.3205
25.5341 | -35.519 | -36.3263
25.5402 | -35.5194 | -36.332
25.5463 | -35.5198 | -36.338
25.5524 | -35.5202 | -36.3439
25.5585 | -35.5205 | -36.3495
25.5646 | -35.5209 | -36.3551
25.5707 | -35.5213 | -36.361
25.5769 | -35.5216 | -36.3667
25.583 | -35.522 | -36.3727
25.5891 | -35.5224 | -36.3785
25.5952 | -35.5227 | -36.3842
25.6013 | -35.5231 | -36.3899
25.6074 | -35.5235 | -36.3955
25.6135 | -35.5239 | -36.4013
25.6196 | -35.5243 | -36.4072
25.6257 | -35.5246 | -36.4129
25.6318 | -35.525 | -36.4186
25.6379 | -35.5253 | -36.4242
25.644 | -35.5257 | -36.4299
25.6501 | -35.5261 | -36.4357
25.6562 | -35.5264 | -36.4415
25.6623 | -35.5268 | -36.4473
25.6684 | -35.5272 | -36.453
25.6745 | -35.5275 | -36.4587
25.6806 | -35.5279 | -36.4642
25.6867 | -35.5283 | -36.4701
25.6928 | -35.5286 | -36.4757
25.6989 | -35.529 | -36.4815
25.705 | -35.5294 | -36.4874
25.7111 | -35.5297 | -36.4932
25.7173 | -35.5301 | -36.4989
25.7234 | -35.5305 | -36.5046
25.7295 | -35.5308 | -36.5104
25.7356 | -35.5312 | -36.5163
25.7417 | -35.5316 | -36.5221
25.7478 | -35.5319 | -36.5277
25.7539 | -35.5323 | -36.5333
25.76 | -35.5327 | -36.5391
25.7661 | -35.533 | -36.5448
25.7722 | -35.5334 | -36.5505
25.7783 | -35.5338 | -36.5562
25.7844 | -35.5342 | -36.5616
25.7905 | -35.5345 | -36.5674
25.7966 | -35.5349 | -36.573
25.8027 | -35.5352 | -36.5788
25.8088 | -35.5356 | -36.5847
25.8149 | -35.536 | -36.5904
25.821 | -35.5363 | -36.596
25.8271 | -35.5367 | -36.6017
25.8332 | -35.537 | -36.6073
25.8393 | -35.5374 | -36.6129
25.8454 | -35.5378 | -36.6186
25.8515 | -35.5382 | -36.6245
25.8576 | -35.5385 | -36.63
25.8638 | -35.5389 | -36.6359
25.8699 | -35.5393 | -36.6417
25.876 | -35.5396 | -36.6475
25.8821 | -35.54 | -36.6535
25.8882 | -35.5403 | -36.6591
25.8943 | -35.5407 | -36.6648
25.9004 | -35.5411 | -36.6706
25.9065 | -35.5415 | -36.6762
25.9126 | -35.5418 | -36.6819
25.9187 | -35.5422 | -36.6877
25.9248 | -35.5425 | -36.6934
25.9309 | -35.5429 | -36.6992
25.937 | -35.5432 | -36.7049
25.9431 | -35.5436 | -36.7106
25.9492 | -35.544 | -36.7165
25.9553 | -35.5443 | -36.722
25.9614 | -35.5447 | -36.7275
25.9675 | -35.5451 | -36.7331
25.9736 | -35.5455 | -36.7386
25.9797 | -35.5458 | -36.7445
25.9858 | -35.5462 | -36.7501
25.9919 | -35.5466 | -36.756
25.998 | -35.547 | -36.7616
26.0042 | -35.5473 | -36.7674
26.0103 | -35.5477 | -36.773
26.0164 | -35.5481 | -36.7788
26.0225 | -35.5484 | -36.7844
26.0286 | -35.5488 | -36.7903
26.0347 | -35.5491 | -36.7959
26.0408 | -35.5495 | -36.8018
26.0469 | -35.5499 | -36.8076
26.053 | -35.5503 | -36.8133
26.0591 | -35.5506 | -36.819
26.0652 | -35.551 | -36.8248
26.0713 | -35.5513 | -36.8306
26.0774 | -35.5517 | -36.8366
26.0835 | -35.5521 | -36.8422
26.0896 | -35.5524 | -36.8478
26.0957 | -35.5528 | -36.8536
26.1018 | -35.5532 | -36.859
26.1079 | -35.5535 | -36.8645
26.114 | -35.5538 | -36.8705
26.1201 | -35.5542 | -36.8764
26.1262 | -35.5545 | -36.8822
26.1323 | -35.5549 | -36.8877
26.1384 | -35.5552 | -36.8933
26.1445 | -35.5556 | -36.8988
26.1507 | -35.556 | -36.9045
26.1568 | -35.5563 | -36.9102
26.1629 | -35.5567 | -36.916
26.169 | -35.557 | -36.9217
26.1751 | -35.5574 | -36.9272
26.1812 | -35.5577 | -36.9329
26.1873 | -35.5581 | -36.9387
26.1934 | -35.5585 | -36.9444
26.1995 | -35.5588 | -36.9501
26.2056 | -35.5592 | -36.9556
26.2117 | -35.5596 | -36.9615
26.2178 | -35.5599 | -36.9671
26.2239 | -35.5603 | -36.9727
26.23 | -35.5606 | -36.9785
26.2361 | -35.561 | -36.9842
26.2422 | -35.5613 | -36.9897
26.2483 | -35.5617 | -36.9953
26.2544 | -35.5621 | -37.001
26.2605 | -35.5625 | -37.0068
26.2666 | -35.5628 | -37.0124
26.2727 | -35.5632 | -37.018
26.2788 | -35.5635 | -37.0235
26.2849 | -35.5638 | -37.0292
26.2911 | -35.5641 | -37.0348
26.2972 | -35.5645 | -37.0405
26.3033 | -35.5649 | -37.046
26.3094 | -35.5652 | -37.0519
26.3155 | -35.5655 | -37.0572
26.3216 | -35.5659 | -37.0629
26.3277 | -35.5663 | -37.0685
26.3338 | -35.5666 | -37.0741
26.3399 | -35.567 | -37.0798
26.346 | -35.5673 | -37.0852
26.3521 | -35.5677 | -37.0909
26.3582 | -35.568 | -37.0966
26.3643 | -35.5684 | -37.1023
26.3704 | -35.5688 | -37.108
26.3765 | -35.5691 | -37.1136
26.3826 | -35.5695 | -37.1193
26.3887 | -35.5699 | -37.1249
26.3948 | -35.5702 | -37.1305
26.4009 | -35.5705 | -37.136
26.407 | -35.5709 | -37.1417
26.4131 | -35.5712 | -37.1472
26.4192 | -35.5716 | -37.1531
26.4253 | -35.572 | -37.1586
26.4314 | -35.5723 | -37.1643
26.4376 | -35.5727 | -37.17
26.4437 | -35.573 | -37.1757
26.4498 | -35.5734 | -37.1813
26.4559 | -35.5737 | -37.1869
26.462 | -35.5741 | -37.1926
26.4681 | -35.5744 | -37.1983
26.4742 | -35.5748 | -37.204
26.4803 | -35.5752 | -37.2096
26.4864 | -35.5755 | -37.2152
26.4925 | -35.5758 | -37.2207
26.4986 | -35.5762 | -37.2264
26.5047 | -35.5765 | -37.232
26.5108 | -35.5769 | -37.2376
26.5169 | -35.5773 | -37.2433
26.523 | -35.5776 | -37.2491
26.5291 | -35.578 | -37.2545
26.5352 | -35.5782 | -37.2604
26.5413 | -35.5786 | -37.2663
26.5474 | -35.579 | -37.2719
26.5535 | -35.5793 | -37.2775
26.5596 | -35.5797 | -37.2832
26.5657 | -35.58 | -37.2888
26.5718 | -35.5804 | -37.2945
26.578 | -35.5807 | -37.3003
26.5841 | -35.5811 | -37.3059
26.5902 | -35.5814 | -37.3115
26.5963 | -35.5818 | -37.3171
26.6024 | -35.5821 | -37.3227
26.6085 | -35.5824 | -37.3284
26.6146 | -35.5828 | -37.334
26.6207 | -35.5831 | -37.3399
26.6268 | -35.5835 | -37.3455
26.6329 | -35.5839 | -37.351
26.639 | -35.5842 | -37.3568
26.6451 | -35.5845 | -37.3623
26.6512 | -35.5849 | -37.3679
26.6573 | -35.5852 | -37.3737
26.6634 | -35.5856 | -37.3793
26.6695 | -35.5859 | -37.3849
26.6756 | -35.5863 | -37.3905
26.6817 | -35.5866 | -37.3961
26.6878 | -35.587 | -37.4018
26.6939 | -35.5873 | -37.4075
26.7 | -35.5877 | -37.4129
26.7061 | -35.588 | -37.4186
26.7122 | -35.5883 | -37.4244
26.7183 | -35.5887 | -37.43
26.7245 | -35.589 | -37.4358
26.7306 | -35.5894 | -37.4414
26.7367 | -35.5897 | -37.447
26.7428 | -35.5901 | -37.4525
26.7489 | -35.5904 | -37.4581
26.755 | -35.5907 | -37.4636
26.7611 | -35.5911 | -37.4692
26.7672 | -35.5914 | -37.4749
26.7733 | -35.5918 | -37.4802
26.7794 | -35.5921 | -37.4861
26.7855 | -35.5925 | -37.4916
26.7916 | -35.5928 | -37.4972
26.7977 | -35.5932 | -37.5028
26.8038 | -35.5935 | -37.5084
26.8099 | -35.5939 | -37.5139
26.816 | -35.5942 | -37.5195
26.8221 | -35.5946 | -37.5252
26.8282 | -35.5949 | -37.531
26.8343 | -35.5952 | -37.5367
26.8404 | -35.5956 | -37.5422
26.8465 | -35.5959 | -37.5478
26.8526 | -35.5963 | -37.5533
26.8587 | -35.5966 | -37.5591
26.8649 | -35.5969 | -37.5645
26.871 | -35.5973 | -37.5702
26.8771 | -35.5977 | -37.5755
26.8832 | -35.598 | -37.5811
26.8893 | -35.5983 | -37.5868
26.8954 | -35.5987 | -37.5923
26.9015 | -35.5991 | -37.598
26.9076 | -35.5994 | -37.6039
26.9137 | -35.5998 | -37.6093
26.9198 | -35.6001 | -37.6149
26.9259 | -35.6005 | -37.6205
26.932 | -35.6008 | -37.626
26.9381 | -35.6012 | -37.6319
26.9442 | -35.6015 | -37.637
26.9503 | -35.6018 | -37.6428
26.9564 | -35.6022 | -37.6486
26.9625 | -35.6026 | -37.6539
26.9686 | -35.6029 | -37.6596
26.9747 | -35.6033 | -37.6653
26.9808 | -35.6036 | -37.6703
26.9869 | -35.604 | -37.6761
26.993 | -35.6043 | -37.6815
26.9991 | -35.6047 | -37.687
27.0052 | -35.605 | -37.6927
27.0114 | -35.6054 | -37.6986
27.0175 | -35.6057 | -37.7041
27.0236 | -35.606 | -37.7095
27.0297 | -35.6064 | -37.7152
27.0358 | -35.6067 | -37.7208
27.0419 | -35.607 | -37.7263
27.048 | -35.6074 | -37.7318
27.0541 | -35.6077 | -37.7373
27.0602 | -35.608 | -37.743
27.0663 | -35.6083 | -37.7484
27.0724 | -35.6086 | -37.7539
27.0785 | -35.609 | -37.7595
27.0846 | -35.6093 | -37.7648
27.0907 | -35.6097 | -37.7704
27.0968 | -35.61 | -37.776
27.1029 | -35.6103 | -37.7815
27.109 | -35.6107 | -37.7871
27.1151 | -35.611 | -37.7925
27.1212 | -35.6113 | -37.7976
27.1273 | -35.6116 | -37.8033
27.1334 | -35.6119 | -37.8089
27.1395 | -35.6123 | -37.8144
27.1456 | -35.6126 | -37.8199
27.1518 | -35.6129 | -37.8256
27.1579 | -35.6132 | -37.8309
27.164 | -35.6135 | -37.8364
27.1701 | -35.6139 | -37.842
27.1762 | -35.6142 | -37.8478
27.1823 | -35.6145 | -37.8535
27.1884 | -35.6149 | -37.859
27.1945 | -35.6152 | -37.8644
27.2006 | -35.6155 | -37.8701
27.2067 | -35.6159 | -37.8756
27.2128 | -35.6162 | -37.8809
27.2189 | -35.6165 | -37.8864
27.225 | -35.6169 | -37.8919
27.2311 | -35.6172 | -37.8974
27.2372 | -35.6175 | -37.9029
27.2433 | -35.6179 | -37.9084
27.2494 | -35.6182 | -37.914
27.2555 | -35.6185 | -37.9195
27.2616 | -35.6188 | -37.925
27.2677 | -35.6192 | -37.9306
27.2738 | -35.6195 | -37.9362
27.2799 | -35.6199 | -37.9415
27.286 | -35.6202 | -37.947
27.2921 | -35.6205 | -37.9525
27.2983 | -35.6208 | -37.9579
27.3044 | -35.6212 | -37.9632
27.3105 | -35.6215 | -37.9688
27.3166 | -35.6218 | -37.9742
27.3227 | -35.6221 | -37.9798
27.3288 | -35.6224 | -37.9853
27.3349 | -35.6227 | -37.9908
27.341 | -35.6231 | -37.9965
27.3471 | -35.6234 | -38.002
27.3532 | -35.6237 | -38.0075
27.3593 | -35.624 | -38.0129
27.3654 | -35.6244 | -38.0185
27.3715 | -35.6247 | -38.0238
27.3776 | -35.625 | -38.0293
27.3837 | -35.6254 | -38.0348
27.3898 | -35.6257 | -38.0406
27.3959 | -35.626 | -38.0462
27.402 | -35.6264 | -38.0515
27.4081 | -35.6267 | -38.0571
27.4142 | -35.627 | -38.0624
27.4203 | -35.6274 | -38.068
27.4264 | -35.6277 | -38.0735
27.4325 | -35.628 | -38.0787
27.4387 | -35.6284 | -38.084
27.4448 | -35.6287 | -38.0898
27.4509 | -35.629 | -38.0953
27.457 | -35.6294 | -38.1006
27.4631 | -35.6297 | -38.1062
27.4692 | -35.63 | -38.1116
27.4753 | -35.6303 | -38.117
27.4814 | -35.6306 | -38.1223
27.4875 | -35.6309 | -38.1279
27.4936 | -35.6313 | -38.1334
27.4997 | -35.6316 | -38.1389
27.5058 | -35.6319 | -38.1443
27.5119 | -35.6322 | -38.1497
27.518 | -35.6325 | -38.1551
27.5241 | -35.6329 | -38.1607
27.5302 | -35.6332 | -38.1663
27.5363 | -35.6336 | -38.1719
27.5424 | -35.6339 | -38.1771
27.5485 | -35.6342 | -38.1825
27.5546 | -35.6346 | -38.1876
27.5607 | -35.6349 | -38.193
27.5668 | -35.6352 | -38.1984
27.5729 | -35.6355 | -38.2038
27.5791 | -35.6358 | -38.2092
27.5852 | -35.6362 | -38.2144
27.5913 | -35.6365 | -38.2198
27.5974 | -35.6368 | -38.2254
27.6035 | -35.6371 | -38.2308
27.6096 | -35.6375 | -38.2364
27.6157 | -35.6377 | -38.2417
27.6218 | -35.6381 | -38.2471
27.6279 | -35.6384 | -38.2526
27.634 | -35.6388 | -38.2583
27.6401 | -35.6391 | -38.2639
27.6462 | -35.6394 | -38.2694
27.6523 | -35.6397 | -38.2745
27.6584 | -35.6401 | -38.28
27.6645 | -35.6404 | -38.2854
27.6706 | -35.6407 | -38.291
27.6767 | -35.641 | -38.2965
27.6828 | -35.6413 | -38.3019
27.6889 | -35.6417 | -38.3074
27.695 | -35.642 | -38.313
27.7011 | -35.6423 | -38.3185
27.7072 | -35.6427 | -38.3239
27.7133 | -35.643 | -38.3295
27.7194 | -35.6433 | -38.3345
27.7256 | -35.6436 | -38.34
27.7317 | -35.6439 | -38.3453
27.7378 | -35.6442 | -38.3511
27.7439 | -35.6446 | -38.3565
27.75 | -35.6449 | -38.3619
27.7561 | -35.6452 | -38.367
27.7622 | -35.6456 | -38.3725
27.7683 | -35.6459 | -38.3782
27.7744 | -35.6462 | -38.3835
27.7805 | -35.6465 | -38.3889
27.7866 | -35.6468 | -38.3943
27.7927 | -35.6471 | -38.3997
27.7988 | -35.6474 | -38.4049
27.8049 | -35.6477 | -38.4103
27.811 | -35.648 | -38.4158
27.8171 | -35.6483 | -38.4211
27.8232 | -35.6486 | -38.4262
27.8293 | -35.6489 | -38.4317
27.8354 | -35.6492 | -38.437
27.8415 | -35.6496 | -38.4424
27.8476 | -35.6499 | -38.4477
27.8537 | -35.6502 | -38.4532
27.8598 | -35.6505 | -38.4586
27.866 | -35.6509 | -38.4638
27.8721 | -35.6512 | -38.469
27.8782 | -35.6515 | -38.4743
27.8843 | -35.6518 | -38.4799
27.8904 | -35.6521 | -38.4852
27.8965 | -35.6525 | -38.4903
27.9026 | -35.6527 | -38.4957
27.9087 | -35.6531 | -38.501
27.9148 | -35.6534 | -38.5063
27.9209 | -35.6537 | -38.512
27.927 | -35.654 | -38.5171
27.9331 | -35.6543 | -38.5224
27.9392 | -35.6546 | -38.5278
27.9453 | -35.6549 | -38.5332
27.9514 | -35.6552 | -38.5387
27.9575 | -35.6556 | -38.5443
27.9636 | -35.6559 | -38.5498
27.9697 | -35.6562 | -38.555
27.9758 | -35.6566 | -38.5605
27.9819 | -35.6569 | -38.5659
27.988 | -35.6572 | -38.5715
27.9941 | -35.6575 | -38.5769
28.0002 | -35.6579 | -38.582
28.0063 | -35.6582 | -38.5876
28.0125 | -35.6584 | -38.5932
28.0186 | -35.6587 | -38.5988
28.0247 | -35.6591 | -38.6038
28.0308 | -35.6594 | -38.6091
28.0369 | -35.6597 | -38.6143
28.043 | -35.66 | -38.6199
28.0491 | -35.6604 | -38.6253
28.0552 | -35.6606 | -38.6308
28.0613 | -35.6609 | -38.6365
28.0674 | -35.6612 | -38.6418
28.0735 | -35.6615 | -38.6469

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5022 | -43.3795
21.4908 | -35.5012 | -35.5148
16.4948 | -34.8856 | -28.0185
12.6602 | -34.5305 | -21.6599
9.7171 | -34.3132 | -16.7043
7.4582 | -34.1682 | -12.6809
5.7244 | -34.056 | -9.4301
4.3936 | -33.998 | -6.7951
3.3722 | -33.9789 | -4.7292
2.5883 | -33.9772 | -3.1593
1.9866 | -33.9732 | -1.907
1.5248 | -33.9663 | -0.90134
1.1703 | -33.9823 | -0.24559
1 | -33.9873 | 0.079469
0.89822 | -33.9836 | 0.35294
0.68941 | -33.9733 | 0.93608
0.52919 | -33.9797 | 1.3899
0.40612 | -33.981 | 2.0039
0.31171 | -33.9791 | 2.4566
0.23928 | -33.9985 | 3.1778
0.18358 | -34.0051 | 3.6033
0.14091 | -33.9995 | 4.4886
0.10819 | -34.0172 | 5.6466
0.083017 | -34.0359 | 7.1149
0.063705 | -34.1041 | 8.9774
0.048923 | -34.178 | 11.4534
0.037479 | -34.2765 | 14.6359
0.028801 | -34.45 | 18.6374
0.022078 | -34.753 | 24.0062
0.017023 | -35.2298 | 30.4634
0.013018 | -36.0629 | 38.1731
0.010014 | -37.1612 | 46.0789

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003628 | -6.2865
0.007257 | -23.6363
0.010885 | -32.4558
0.014513 | -40.168
0.018141 | -43.2982
0.02177 | -44.7974
0.025398 | -46.2407
0.029026 | -47.7958
0.032655 | -48.6075
0.036283 | -49.6871
0.039911 | -51.6949
0.043539 | -52.8934
0.047168 | -54.3603
0.050796 | -54.9555
0.054424 | -56.8626
0.058053 | -56.7901
0.061681 | -57.559
0.065309 | -57.0903
0.068937 | -57.3529
0.072566 | -57.8223
0.076194 | -59.5721
0.079822 | -60.7073
0.08345 | -60.9326
0.087079 | -60.6472
0.090707 | -60.8868
0.094335 | -62.3732
0.097964 | -62.6999
0.10159 | -62.1609
0.10522 | -62.6735
0.10885 | -63.4759
0.11248 | -63.8277
0.1161 | -64.1923
0.11973 | -64.4814
0.12336 | -64.9849
0.12699 | -65.5682
0.13062 | -66.8002
0.13425 | -67.8129
0.13787 | -66.479
0.1415 | -66.505
0.14513 | -66.6497
0.14876 | -66.735
0.15239 | -66.6515
0.15602 | -68.1586
0.15964 | -69.3351
0.16327 | -69.6747
0.1669 | -69.0595
0.17053 | -69.4987
0.17416 | -69.2912
0.17779 | -69.6945
0.18141 | -69.1946
0.18504 | -69.2228
0.18867 | -69.3041
0.1923 | -69.975
0.19593 | -70.0695
0.19956 | -70.5056
0.20318 | -72.0204
0.20681 | -73.2162
0.21044 | -72.0354
0.21407 | -71.5809
0.2177 | -71.4314
0.22132 | -71.1794
0.22495 | -72.2719
0.22858 | -73.7453
0.23221 | -72.708
0.23584 | -72.6856
0.23947 | -73.1789
0.2431 | -72.5364
0.24672 | -73.0047
0.25035 | -72.5836
0.25398 | -73.0289
0.25761 | -74.366
0.26124 | -75.2643
0.26487 | -74.3586
0.26849 | -74.2416
0.27212 | -74.6886
0.27575 | -74.7615
0.27938 | -75.0448
0.28301 | -76.3492
0.28663 | -75.682
0.29026 | -74.7259
0.29389 | -74.2163
0.29752 | -75.0468
0.30115 | -75.5398
0.30478 | -76.2045
0.3084 | -74.3533
0.31203 | -74.8712
0.31566 | -75.6314
0.31929 | -75.6639
0.32292 | -76.8292
0.32654 | -76.8527
0.33017 | -76.516
0.3338 | -75.9114
0.33743 | -75.927
0.34106 | -76.8939
0.34469 | -77.4392
0.34831 | -77.4759
0.35194 | -77.433
0.35557 | -76.1156
0.3592 | -76.5839
0.36283 | -76.8014
0.36646 | -77.0653
0.37008 | -77.1234
0.37371 | -78.5654
0.37734 | -78.5795
0.38097 | -77.7645
0.3846 | -77.0816
0.38823 | -77.19
0.39186 | -78.3894
0.39548 | -79.9847
0.39911 | -79.9896
0.40274 | -79.1107
0.40637 | -79.4926
0.41 | -78.7762
0.41362 | -78.7241
0.41725 | -79.9429
0.42088 | -79.6971
0.42451 | -78.7652
0.42814 | -78.1141
0.43177 | -78.2482
0.43539 | -78.7409
0.43902 | -78.2714
0.44265 | -77.6541
0.44628 | -79.1253
0.44991 | -79.4801
0.45354 | -79.2632
0.45716 | -79.6563
0.46079 | -79.6719
0.46442 | -78.7774
0.46805 | -79.3775
0.47168 | -80.0599
0.4753 | -80.1286
0.47893 | -80.3077
0.48256 | -79.816
0.48619 | -79.6463
0.48982 | -79.4373
0.49345 | -80.0218
0.49707 | -79.5473
0.5007 | -80.1846
0.50433 | -81.2341
0.50796 | -81.6008
0.51159 | -81.6455
0.51522 | -80.8893
0.51884 | -80.6643
0.52247 | -80.6158
0.5261 | -80.939
0.52973 | -81.75
0.53336 | -82.4335
0.53699 | -82.1934
0.54061 | -81.3874
0.54424 | -81.1383
0.54787 | -81.0134
0.5515 | -80.8919
0.55513 | -80.9383
0.55876 | -80.5669
0.56238 | -80.4796
0.56601 | -81.143
0.56964 | -81.8499
0.57327 | -82.0006
0.5769 | -81.8908
0.58052 | -81.4513
0.58415 | -80.9109
0.58778 | -80.5808
0.59141 | -80.7838
0.59504 | -81.3677
0.59867 | -82.0087
0.6023 | -82.4982
0.60592 | -82.6652
0.60955 | -82.5236
0.61318 | -82.3538
0.61681 | -82.3
0.62044 | -82.3853
0.62406 | -82.5341
0.62769 | -82.8414
0.63132 | -83.522
0.63495 | -83.7948
0.63858 | -83.3326
0.64221 | -83.0914
0.64583 | -82.9221
0.64946 | -82.4999
0.65309 | -82.4637
0.65672 | -83.1721
0.66035 | -83.6794
0.66398 | -83.5757
0.6676 | -83.5768
0.67123 | -83.3475
0.67486 | -83.1595
0.67849 | -83.1996
0.68212 | -83.1676
0.68575 | -83.3953
0.68937 | -83.755
0.693 | -83.7611
0.69663 | -83.6578
0.70026 | -83.5933
0.70389 | -83.5344
0.70752 | -83.5792
0.71114 | -83.8621
0.71477 | -84.1239
0.7184 | -84.5875
0.72203 | -84.8016
0.72566 | -84.3581
0.72928 | -83.9435
0.73291 | -83.2911
0.73654 | -82.8327
0.74017 | -83.1423
0.7438 | -83.7802
0.74743 | -84.2143
0.75105 | -83.9237
0.75468 | -83.864
0.75831 | -84.3421
0.76194 | -84.5596
0.76557 | -84.4536
0.7692 | -84.4271
0.77282 | -84.4707
0.77645 | -84.6735
0.78008 | -84.6207
0.78371 | -84.0857
0.78734 | -83.6329
0.79097 | -83.3086
0.79459 | -82.9911
0.79822 | -83.2314
0.80185 | -83.9627
0.80548 | -84.4396
0.80911 | -84.7678
0.81273 | -85.11
0.81636 | -84.6975
0.81999 | -83.8534
0.82362 | -83.7456
0.82725 | -84.1489
0.83088 | -84.7578
0.83451 | -85.3628
0.83813 | -85.6242
0.84176 | -85.5366
0.84539 | -85.419
0.84902 | -85.6101
0.85265 | -85.9358
0.85628 | -86.0897
0.8599 | -85.9306
0.86353 | -85.5583
0.86716 | -85.2118
0.87079 | -85.3242
0.87442 | -85.9347
0.87804 | -85.8827
0.88167 | -85.5695
0.8853 | -85.3562
0.88893 | -85.2629
0.89256 | -85.104
0.89619 | -85.2431
0.89981 | -85.577
0.90344 | -85.4856
0.90707 | -85.3567
0.9107 | -85.6763
0.91433 | -85.9748
0.91795 | -85.8629
0.92158 | -85.9645
0.92521 | -86.0402
0.92884 | -86.147
0.93247 | -86.4355
0.9361 | -86.5563
0.93973 | -86.6034
0.94335 | -86.7898
0.94698 | -86.884
0.95061 | -86.9252
0.95424 | -86.905
0.95787 | -86.3675
0.96149 | -85.756
0.96512 | -85.8606
0.96875 | -86.2176
0.97238 | -86.5246
0.97601 | -86.4987
0.97964 | -86.4195
0.98326 | -86.5806
0.98689 | -86.835
0.99052 | -86.7433
0.99415 | -86.3535
0.99778 | -86.3002
1.0014 | -86.4687
1.005 | -86.9061
1.0087 | -87.4195
1.0123 | -87.3611
1.0159 | -87.26
1.0195 | -87.5321
1.0232 | -87.4952
1.0268 | -87.0871
1.0304 | -86.8848
1.0341 | -86.9773
1.0377 | -87.1809
1.0413 | -87.366
1.0449 | -87.5992
1.0486 | -87.8941
1.0522 | -88.1979
1.0558 | -88.2778
1.0595 | -88.0206
1.0631 | -87.5922
1.0667 | -87.1754
1.0703 | -86.9427
1.074 | -87.0306
1.0776 | -87.3955
1.0812 | -87.8352
1.0849 | -88.1348
1.0885 | -88.2193
1.0921 | -88.1733
1.0957 | -88.0819
1.0994 | -87.9666
1.103 | -87.8828
1.1066 | -87.9698
1.1103 | -88.2964
1.1139 | -88.6261
1.1175 | -88.5827
1.1211 | -88.1585
1.1248 | -87.7103
1.1284 | -87.5553
1.132 | -87.7213
1.1357 | -87.9663
1.1393 | -88.0767
1.1429 | -88.0633
1.1465 | -88.0243
1.1502 | -87.9993
1.1538 | -88.059
1.1574 | -88.258
1.161 | -88.4586
1.1647 | -88.5167
1.1683 | -88.5842
1.1719 | -88.7993
1.1756 | -88.9851
1.1792 | -88.9454
1.1828 | -88.7941
1.1864 | -88.6412
1.1901 | -88.4672
1.1937 | -88.3248
1.1973 | -88.302
1.201 | -88.3509
1.2046 | -88.4826
1.2082 | -88.6882
1.2118 | -88.7338
1.2155 | -88.5403
1.2191 | -88.3656
1.2227 | -88.3867
1.2264 | -88.6034
1.23 | -88.9043
1.2336 | -89.0862
1.2372 | -89.0734
1.2409 | -89.055
1.2445 | -89.1274
1.2481 | -89.2148
1.2518 | -89.2641
1.2554 | -89.3274
1.259 | -89.4183
1.2626 | -89.4321
1.2663 | -89.3867
1.2699 | -89.3877
1.2735 | -89.4148
1.2772 | -89.4525
1.2808 | -89.4593
1.2844 | -89.3811
1.288 | -89.3016
1.2917 | -89.3213
1.2953 | -89.416
1.2989 | -89.3825
1.3026 | -89.0757
1.3062 | -88.675
1.3098 | -88.4694
1.3134 | -88.6155
1.3171 | -89.1048
1.3207 | -89.6862
1.3243 | -89.9853
1.328 | -89.9215
1.3316 | -89.7501
1.3352 | -89.6598
1.3388 | -89.6186
1.3425 | -89.5835
1.3461 | -89.6007
1.3497 | -89.5652
1.3533 | -89.4319
1.357 | -89.4746
1.3606 | -89.7878
1.3642 | -90.0304
1.3679 | -89.9979
1.3715 | -89.8749
1.3751 | -89.8734
1.3787 | -90.0664
1.3824 | -90.2613
1.386 | -90.2053
1.3896 | -89.9705
1.3933 | -89.7753
1.3969 | -89.7331
1.4005 | -89.8227
1.4041 | -89.9625
1.4078 | -90.1111
1.4114 | -90.2455
1.415 | -90.3498
1.4187 | -90.4893
1.4223 | -90.6414
1.4259 | -90.6247
1.4295 | -90.3523
1.4332 | -89.9204
1.4368 | -89.6135
1.4404 | -89.7057
1.4441 | -90.1469
1.4477 | -90.5733
1.4513 | -90.6137
1.4549 | -90.3676
1.4586 | -90.2532
1.4622 | -90.2692
1.4658 | -90.0747
1.4695 | -89.7758
1.4731 | -89.7805
1.4767 | -90.1226
1.4803 | -90.52
1.484 | -90.7607
1.4876 | -90.8318
1.4912 | -90.8843
1.4949 | -90.9954
1.4985 | -90.9617
1.5021 | -90.7282
1.5057 | -90.6269
1.5094 | -90.8587
1.513 | -91.2385
1.5166 | -91.3978
1.5203 | -91.1044
1.5239 | -90.5384
1.5275 | -90.1456
1.5311 | -90.1208
1.5348 | -90.2194
1.5384 | -90.1594
1.542 | -90.0503
1.5456 | -90.0571
1.5493 | -90.1131
1.5529 | -90.2246
1.5565 | -90.4797
1.5602 | -90.7342
1.5638 | -90.7628
1.5674 | -90.7423
1.571 | -90.8817
1.5747 | -91.1941
1.5783 | -91.6316
1.5819 | -92.0077
1.5856 | -92.1892
1.5892 | -92.2087
1.5928 | -92.1036
1.5964 | -91.8064
1.6001 | -91.2773
1.6037 | -90.6829
1.6073 | -90.2543
1.611 | -90.1021
1.6146 | -90.2127
1.6182 | -90.5419
1.6218 | -91.0295
1.6255 | -91.5482
1.6291 | -91.8904
1.6327 | -91.9638
1.6364 | -91.8154
1.64 | -91.5814
1.6436 | -91.3984
1.6472 | -91.3529
1.6509 | -91.4422
1.6545 | -91.6012
1.6581 | -91.7754
1.6618 | -91.9658
1.6654 | -92.1514
1.669 | -92.2525
1.6726 | -92.2092
1.6763 | -92.0586
1.6799 | -91.9102
1.6835 | -91.8563
1.6872 | -91.9171
1.6908 | -92.0102
1.6944 | -91.9857
1.698 | -91.7795
1.7017 | -91.4831
1.7053 | -91.2598
1.7089 | -91.202
1.7125 | -91.2929
1.7162 | -91.4366
1.7198 | -91.5158
1.7234 | -91.5137
1.7271 | -91.5105
1.7307 | -91.5266
1.7343 | -91.5162
1.7379 | -91.4693
1.7416 | -91.4717
1.7452 | -91.5515
1.7488 | -91.6568
1.7525 | -91.7468
1.7561 | -91.8285
1.7597 | -91.8935
1.7633 | -91.8924
1.767 | -91.7906
1.7706 | -91.6439
1.7742 | -91.5468
1.7779 | -91.5824
1.7815 | -91.7789
1.7851 | -92.0658
1.7887 | -92.282
1.7924 | -92.2789
1.796 | -92.061
1.7996 | -91.8251
1.8033 | -91.7316
1.8069 | -91.7766
1.8105 | -91.8694
1.8141 | -91.9539
1.8178 | -92.0556
1.8214 | -92.2301
1.825 | -92.4659
1.8287 | -92.6559
1.8323 | -92.6672
1.8359 | -92.4992
1.8395 | -92.3045
1.8432 | -92.2131
1.8468 | -92.222
1.8504 | -92.2209
1.8541 | -92.1398
1.8577 | -92.0218
1.8613 | -91.9325
1.8649 | -91.9177
1.8686 | -91.9674
1.8722 | -92.0065
1.8758 | -91.9902
1.8795 | -91.9879
1.8831 | -92.061
1.8867 | -92.1616
1.8903 | -92.2624
1.894 | -92.379
1.8976 | -92.502
1.9012 | -92.5716
1.9048 | -92.5434
1.9085 | -92.4907
1.9121 | -92.5349
1.9157 | -92.7307
1.9194 | -92.9641
1.923 | -93.0231
1.9266 | -92.8446
1.9302 | -92.5372
1.9339 | -92.2609
1.9375 | -92.126
1.9411 | -92.1644
1.9448 | -92.2829
1.9484 | -92.3645
1.952 | -92.4252
1.9556 | -92.5452
1.9593 | -92.7512
1.9629 | -93.0219
1.9665 | -93.2973
1.9702 | -93.4774
1.9738 | -93.4858
1.9774 | -93.3574
1.981 | -93.1834
1.9847 | -93.0222
1.9883 | -92.9203
1.9919 | -92.8827
1.9956 | -92.8875
1.9992 | -92.9206
2.0028 | -93.0065
2.0064 | -93.1868
2.0101 | -93.4059
2.0137 | -93.4951
2.0173 | -93.3642
2.021 | -93.1243
2.0246 | -92.9413
2.0282 | -92.8757
2.0318 | -92.9251
2.0355 | -93.0525
2.0391 | -93.1485
2.0427 | -93.1393
2.0464 | -93.0735
2.05 | -93.0424
2.0536 | -93.0646
2.0572 | -93.0715
2.0609 | -93.0169
2.0645 | -92.965
2.0681 | -92.9981
2.0717 | -93.1517
2.0754 | -93.3555
2.079 | -93.482
2.0826 | -93.4644
2.0863 | -93.3545
2.0899 | -93.2702
2.0935 | -93.292
2.0971 | -93.318
2.1008 | -93.368
2.1044 | -93.4416
2.108 | -93.5162
2.1117 | -93.5574
2.1153 | -93.5403
2.1189 | -93.49
2.1225 | -93.4574
2.1262 | -93.5096
2.1298 | -93.6411
2.1334 | -93.7846
2.1371 | -93.8876
2.1407 | -93.9392
2.1443 | -93.9651
2.1479 | -93.9719
2.1516 | -93.9382
2.1552 | -93.8546
2.1588 | -93.7248
2.1625 | -93.5857
2.1661 | -93.4561
2.1697 | -93.3661
2.1733 | -93.3492
2.177 | -93.4425
2.1806 | -93.637
2.1842 | -93.8704
2.1879 | -94.0747
2.1915 | -94.214
2.1951 | -94.2647
2.1987 | -94.22
2.2024 | -94.0985
2.206 | -93.9744
2.2096 | -93.8981
2.2133 | -93.8716
2.2169 | -93.8739
2.2205 | -93.91
2.2241 | -94.0119
2.2278 | -94.1968
2.2314 | -94.4349
2.235 | -94.6349
2.2386 | -94.7158
2.2423 | -94.6229
2.2459 | -94.3861
2.2495 | -94.1206
2.2532 | -93.9389
2.2568 | -93.8956
2.2604 | -93.9728
2.264 | -94.0938
2.2677 | -94.1902
2.2713 | -94.2335
2.2749 | -94.2423
2.2786 | -94.2153
2.2822 | -94.1766
2.2858 | -94.152
2.2894 | -94.1751
2.2931 | -94.2717
2.2967 | -94.4245
2.3003 | -94.5685
2.304 | -94.6189
2.3076 | -94.5574
2.3112 | -94.4153
2.3148 | -94.254
2.3185 | -94.1423
2.3221 | -94.1067
2.3257 | -94.1491
2.3294 | -94.2463
2.333 | -94.3435
2.3366 | -94.3959
2.3402 | -94.3695
2.3439 | -94.2949
2.3475 | -94.2028
2.3511 | -94.1198
2.3548 | -94.0382
2.3584 | -93.9508
2.362 | -93.8848
2.3656 | -93.8927
2.3693 | -94.0022
2.3729 | -94.1992
2.3765 | -94.4102
2.3802 | -94.5486
2.3838 | -94.546
2.3874 | -94.449
2.391 | -94.3353
2.3947 | -94.264
2.3983 | -94.2665
2.4019 | -94.3328
2.4056 | -94.4113
2.4092 | -94.4594
2.4128 | -94.4775
2.4164 | -94.478
2.4201 | -94.4726
2.4237 | -94.4693
2.4273 | -94.4646
2.4309 | -94.4832
2.4346 | -94.5345
2.4382 | -94.6009
2.4418 | -94.6535
2.4455 | -94.6725
2.4491 | -94.664
2.4527 | -94.6436
2.4563 | -94.6429
2.46 | -94.6726
2.4636 | -94.7058
2.4672 | -94.725
2.4709 | -94.7441
2.4745 | -94.8
2.4781 | -94.9071
2.4817 | -95.0356
2.4854 | -95.1195
2.489 | -95.0881
2.4926 | -94.9392
2.4963 | -94.7507
2.4999 | -94.6029
2.5035 | -94.5632
2.5071 | -94.6238
2.5108 | -94.7257
2.5144 | -94.8071
2.518 | -94.8589
2.5217 | -94.8825
2.5253 | -94.8819
2.5289 | -94.8633
2.5325 | -94.814
2.5362 | -94.7344
2.5398 | -94.6602
2.5434 | -94.6194
2.5471 | -94.635
2.5507 | -94.6833
2.5543 | -94.7127
2.5579 | -94.6851
2.5616 | -94.6244
2.5652 | -94.5772
2.5688 | -94.5712
2.5725 | -94.6519
2.5761 | -94.8396
2.5797 | -95.1009
2.5833 | -95.3553
2.587 | -95.5082
2.5906 | -95.544
2.5942 | -95.4886
2.5979 | -95.355
2.6015 | -95.179
2.6051 | -94.9862
2.6087 | -94.8216
2.6124 | -94.7143
2.616 | -94.7834
2.6196 | -94.866
2.6232 | -95.0297
2.6269 | -95.2542
2.6305 | -95.4972
2.6341 | -95.6972
2.6378 | -95.8183
2.6414 | -95.8423
2.645 | -95.7745
2.6486 | -95.6612
2.6523 | -95.5406
2.6559 | -95.4493
2.6595 | -95.4096
2.6632 | -95.4172
2.6668 | -95.4605
2.6704 | -95.5216
2.674 | -95.5774
2.6777 | -95.6225
2.6813 | -95.6479
2.6849 | -95.6578
2.6886 | -95.6626
2.6922 | -95.6739
2.6958 | -95.7122
2.6994 | -95.7792
2.7031 | -95.8667
2.7067 | -95.9625
2.7103 | -96.0331
2.714 | -96.0539
2.7176 | -96.0232
2.7212 | -95.9404
2.7248 | -95.8123
2.7285 | -95.6841
2.7321 | -95.5654
2.7357 | -95.4697
2.7394 | -95.3735
2.743 | -95.2677
2.7466 | -95.1663
2.7502 | -95.0751
2.7539 | -95.002
2.7575 | -94.9548
2.7611 | -94.9248
2.7648 | -94.9033
2.7684 | -94.916
2.772 | -94.9874
2.7756 | -95.1206
2.7793 | -95.3118
2.7829 | -95.512
2.7865 | -95.6806
2.7901 | -95.7833
2.7938 | -95.7977
2.7974 | -95.7338
2.801 | -95.6132
2.8047 | -95.4785
2.8083 | -95.3906
2.8119 | -95.3881
2.8155 | -95.4399
2.8192 | -95.5032
2.8228 | -95.5369
2.8264 | -95.5271
2.8301 | -95.4947
2.8337 | -95.4803
2.8373 | -95.4882
2.8409 | -95.5248
2.8446 | -95.5834
2.8482 | -95.6423
2.8518 | -95.679
2.8555 | -95.673
2.8591 | -95.6368
2.8627 | -95.5954
2.8663 | -95.5705
2.87 | -95.5789
2.8736 | -95.6311
2.8772 | -95.7294
2.8809 | -95.8524
2.8845 | -95.9893
2.8881 | -96.0925
2.8917 | -96.1248
2.8954 | -96.0906
2.899 | -96.0083
2.9026 | -95.9259
2.9063 | -95.869
2.9099 | -95.8692
2.9135 | -95.9138
2.9171 | -95.9701
2.9208 | -96.0054
2.9244 | -96.0061
2.928 | -96.0007
2.9317 | -96.0168
2.9353 | -96.0842
2.9389 | -96.2057
2.9425 | -96.3497
2.9462 | -96.4792
2.9498 | -96.5621
2.9534 | -96.5973
2.9571 | -96.5884
2.9607 | -96.5374
2.9643 | -96.4621
2.9679 | -96.3882
2.9716 | -96.3328
2.9752 | -96.3165
2.9788 | -96.3485
2.9824 | -96.4021
2.9861 | -96.433
2.9897 | -96.4216
2.9933 | -96.3664
2.997 | -96.2861
3.0006 | -96.2131
3.0042 | -96.1679
3.0078 | -96.1702
3.0115 | -96.1901
3.0151 | -96.2086
3.0187 | -96.2188
3.0224 | -96.205
3.026 | -96.1633
3.0296 | -96.1122
3.0332 | -96.066
3.0369 | -96.0298
3.0405 | -96.0028
3.0441 | -95.9836
3.0478 | -95.9645
3.0514 | -95.951
3.055 | -95.9391
3.0586 | -95.9463
3.0623 | -95.9711
3.0659 | -96.0254
3.0695 | -96.1193
3.0732 | -96.2452
3.0768 | -96.3643
3.0804 | -96.4392
3.084 | -96.4539
3.0877 | -96.4305
3.0913 | -96.39
3.0949 | -96.3427
3.0986 | -96.2867
3.1022 | -96.2463
3.1058 | -96.2394
3.1094 | -96.298
3.1131 | -96.419
3.1167 | -96.5651
3.1203 | -96.6722
3.124 | -96.6979
3.1276 | -96.6434
3.1312 | -96.5321
3.1348 | -96.3988
3.1385 | -96.3267
3.1421 | -96.2558
3.1457 | -96.2118
3.1493 | -96.1846
3.153 | -96.1561
3.1566 | -96.1301
3.1602 | -96.1023
3.1639 | -96.0859
3.1675 | -96.1059
3.1711 | -96.1749
3.1747 | -96.2857
3.1784 | -96.4203
3.182 | -96.5742
3.1856 | -96.7372
3.1893 | -96.9016
3.1929 | -97.0536
3.1965 | -97.1974
3.2001 | -97.3254
3.2038 | -97.4302
3.2074 | -97.5088
3.211 | -97.5443
3.2147 | -97.5231
3.2183 | -97.4249
3.2219 | -97.245
3.2255 | -97.0296
3.2292 | -96.8147
3.2328 | -96.6381
3.2364 | -96.5151
3.2401 | -96.463
3.2437 | -96.4681
3.2473 | -96.5129
3.2509 | -96.5672
3.2546 | -96.6226
3.2582 | -96.6665
3.2618 | -96.6899
3.2655 | -96.6911
3.2691 | -96.691
3.2727 | -96.7036
3.2763 | -96.729
3.28 | -96.7545
3.2836 | -96.7818
3.2872 | -96.8018
3.2909 | -96.8021
3.2945 | -96.789
3.2981 | -96.7774
3.3017 | -96.7659
3.3054 | -96.7589
3.309 | -96.752
3.3126 | -96.7553
3.3163 | -96.7599
3.3199 | -96.7661
3.3235 | -96.7586
3.3271 | -96.7374
3.3308 | -96.7307
3.3344 | -96.7539
3.338 | -96.8215
3.3416 | -96.9377
3.3453 | -97.0906
3.3489 | -97.249
3.3525 | -97.392
3.3562 | -97.5007
3.3598 | -97.5664
3.3634 | -97.5739
3.367 | -97.5339
3.3707 | -97.4639
3.3743 | -97.3673
3.3779 | -97.2715
3.3816 | -97.2024
3.3852 | -97.1505
3.3888 | -97.1068
3.3924 | -97.0625
3.3961 | -97.0137
3.3997 | -96.9665
3.4033 | -96.9227
3.407 | -96.9005
3.4106 | -96.9062
3.4142 | -96.9383
3.4178 | -97.0085
3.4215 | -97.107
3.4251 | -97.2399
3.4287 | -97.4032
3.4324 | -97.5826
3.436 | -97.7603
3.4396 | -97.9386
3.4432 | -98.1028
3.4469 | -98.2203
3.4505 | -98.2689
3.4541 | -98.2544
3.4578 | -98.1853
3.4614 | -98.0672
3.465 | -97.9328
3.4686 | -97.7762
3.4723 | -97.6011
3.4759 | -97.4245
3.4795 | -97.2591
3.4832 | -97.1319
3.4868 | -97.0431
3.4904 | -96.9936
3.494 | -96.9842
3.4977 | -97.0076
3.5013 | -97.0475
3.5049 | -97.0942
3.5085 | -97.1393
3.5122 | -97.1837
3.5158 | -97.2214
3.5194 | -97.2309
3.5231 | -97.2284
3.5267 | -97.2138
3.5303 | -97.2053
3.5339 | -97.2245
3.5376 | -97.2835
3.5412 | -97.3874
3.5448 | -97.5166
3.5485 | -97.6457
3.5521 | -97.7534
3.5557 | -97.8285
3.5593 | -97.8611
3.563 | -97.8394
3.5666 | -97.7847
3.5702 | -97.7236
3.5739 | -97.6771
3.5775 | -97.6528
3.5811 | -97.6622
3.5847 | -97.6867
3.5884 | -97.7078
3.592 | -97.7137
3.5956 | -97.7024
3.5993 | -97.6775
3.6029 | -97.6346
3.6065 | -97.5913
3.6101 | -97.5303
3.6138 | -97.4568
3.6174 | -97.3886
3.621 | -97.336
3.6247 | -97.3147
3.6283 | -97.3302
3.6319 | -97.3697
3.6355 | -97.422
3.6392 | -97.4795
3.6428 | -97.5232
3.6464 | -97.5537
3.6501 | -97.5709
3.6537 | -97.5894
3.6573 | -97.6108
3.6609 | -97.6157
3.6646 | -97.6144
3.6682 | -97.6116
3.6718 | -97.6021
3.6755 | -97.5926
3.6791 | -97.5853
3.6827 | -97.5875
3.6863 | -97.6018
3.69 | -97.6174
3.6936 | -97.6258
3.6972 | -97.6223
3.7008 | -97.6092
3.7045 | -97.5821
3.7081 | -97.5603
3.7117 | -97.5424
3.7154 | -97.5397
3.719 | -97.57
3.7226 | -97.6201
3.7262 | -97.6896
3.7299 | -97.7683
3.7335 | -97.8451
3.7371 | -97.9018
3.7408 | -97.9345
3.7444 | -97.9306
3.748 | -97.889
3.7516 | -97.8214
3.7553 | -97.7377
3.7589 | -97.6544
3.7625 | -97.5924
3.7662 | -97.5748
3.7698 | -97.6015
3.7734 | -97.681
3.777 | -97.8119
3.7807 | -97.9625
3.7843 | -98.1042
3.7879 | -98.2179
3.7916 | -98.2841
3.7952 | -98.2894
3.7988 | -98.2595
3.8024 | -98.195
3.8061 | -98.106
3.8097 | -98.0153
3.8133 | -97.9348
3.817 | -97.8739
3.8206 | -97.8385
3.8242 | -97.8273
3.8278 | -97.8244
3.8315 | -97.8359
3.8351 | -97.8537
3.8387 | -97.8637
3.8424 | -97.8699
3.846 | -97.8781
3.8496 | -97.8814
3.8532 | -97.881
3.8569 | -97.8891
3.8605 | -97.9047
3.8641 | -97.9438
3.8677 | -97.9924
3.8714 | -98.0389
3.875 | -98.0751
3.8786 | -98.1029
3.8823 | -98.1148
3.8859 | -98.1154
3.8895 | -98.1156
3.8931 | -98.121
3.8968 | -98.1342
3.9004 | -98.1569
3.904 | -98.1918
3.9077 | -98.222
3.9113 | -98.2369
3.9149 | -98.245
3.9185 | -98.2367
3.9222 | -98.2089
3.9258 | -98.1677
3.9294 | -98.1272
3.9331 | -98.068
3.9367 | -98.0087
3.9403 | -97.9477
3.9439 | -97.9059
3.9476 | -97.8821
3.9512 | -97.8833
3.9548 | -97.9018
3.9585 | -97.9454
3.9621 | -98.0045
3.9657 | -98.0627
3.9693 | -98.1278
3.973 | -98.2026
3.9766 | -98.2773
3.9802 | -98.3612
3.9839 | -98.4641
3.9875 | -98.5793
3.9911 | -98.6971
3.9947 | -98.8142
3.9984 | -98.907
4.002 | -98.975
4.0056 | -98.9922
4.0093 | -98.9673
4.0129 | -98.9031
4.0165 | -98.8132
4.0201 | -98.7157
4.0238 | -98.6341
4.0274 | -98.5769
4.031 | -98.5323
4.0347 | -98.5206
4.0383 | -98.5399
4.0419 | -98.5823
4.0455 | -98.6348
4.0492 | -98.7051
4.0528 | -98.789
4.0564 | -98.8844
4.06 | -98.9749
4.0637 | -99.0423
4.0673 | -99.0797
4.0709 | -99.0606
4.0746 | -98.9895
4.0782 | -98.8822
4.0818 | -98.7466
4.0854 | -98.6063
4.0891 | -98.4821
4.0927 | -98.382
4.0963 | -98.3006
4.1 | -98.2588
4.1036 | -98.2531
4.1072 | -98.2828
4.1108 | -98.3342
4.1145 | -98.3963
4.1181 | -98.4586
4.1217 | -98.512
4.1254 | -98.5513
4.129 | -98.5728
4.1326 | -98.5829
4.1362 | -98.5812
4.1399 | -98.5848
4.1435 | -98.6141
4.1471 | -98.6745
4.1508 | -98.7621
4.1544 | -98.8664
4.158 | -98.9578
4.1616 | -99.0229
4.1653 | -99.0422
4.1689 | -99.0091
4.1725 | -98.9409
4.1762 | -98.8529
4.1798 | -98.7607
4.1834 | -98.6943
4.187 | -98.6907
4.1907 | -98.6706
4.1943 | -98.6635
4.1979 | -98.6647
4.2016 | -98.6723
4.2052 | -98.6836
4.2088 | -98.6927
4.2124 | -98.6913
4.2161 | -98.6846
4.2197 | -98.6734
4.2233 | -98.6633
4.2269 | -98.6503
4.2306 | -98.6479
4.2342 | -98.6564
4.2378 | -98.6778
4.2415 | -98.7175
4.2451 | -98.7768
4.2487 | -98.847
4.2523 | -98.9323
4.256 | -99.0303
4.2596 | -99.1301
4.2632 | -99.2105
4.2669 | -99.2517
4.2705 | -99.2503
4.2741 | -99.1992
4.2777 | -99.1071
4.2814 | -98.9931
4.285 | -98.8669
4.2886 | -98.7545
4.2923 | -98.6612
4.2959 | -98.603
4.2995 | -98.5788
4.3031 | -98.5883
4.3068 | -98.6247
4.3104 | -98.6769
4.314 | -98.7364
4.3177 | -98.7967
4.3213 | -98.8523
4.3249 | -98.8929
4.3285 | -98.9243
4.3322 | -98.9517
4.3358 | -98.977
4.3394 | -98.9988
4.3431 | -99.0247
4.3467 | -99.0533
4.3503 | -99.0707
4.3539 | -99.0722
4.3576 | -99.0541
4.3612 | -99.0164
4.3648 | -98.9607
4.3685 | -98.9036
4.3721 | -98.8502
4.3757 | -98.8119
4.3793 | -98.7962
4.383 | -98.8012
4.3866 | -98.8273
4.3902 | -98.8754
4.3938 | -98.951
4.3975 | -99.0441
4.4011 | -99.1597
4.4047 | -99.2935
4.4084 | -99.431
4.412 | -99.5563
4.4156 | -99.6668
4.4192 | -99.7567
4.4229 | -99.8047
4.4265 | -99.8107
4.4301 | -99.7846
4.4338 | -99.7283
4.4374 | -99.6479
4.441 | -99.5633
4.4446 | -99.4851
4.4483 | -99.428
4.4519 | -99.4004
4.4555 | -99.4086
4.4592 | -99.4528
4.4628 | -99.5279
4.4664 | -99.6287
4.47 | -99.7405
4.4737 | -99.8459
4.4773 | -99.9451
4.4809 | -100.0286
4.4846 | -100.0806
4.4882 | -100.1051
4.4918 | -100.1106
4.4954 | -100.0897
4.4991 | -100.0421
4.5027 | -99.9741
4.5063 | -99.8967
4.51 | -99.8215
4.5136 | -99.7615
4.5172 | -99.7218
4.5208 | -99.7094
4.5245 | -99.719
4.5281 | -99.7406
4.5317 | -99.7733
4.5354 | -99.8143
4.539 | -99.8494
4.5426 | -99.8749
4.5462 | -99.8801
4.5499 | -99.8782
4.5535 | -99.8662
4.5571 | -99.8367
4.5608 | -99.8071
4.5644 | -99.7835
4.568 | -99.7706
4.5716 | -99.7668
4.5753 | -99.7558
4.5789 | -99.7359
4.5825 | -99.7009
4.5861 | -99.6579
4.5898 | -99.5973
4.5934 | -99.5269
4.597 | -99.4551
4.6007 | -99.3842
4.6043 | -99.3055
4.6079 | -99.2316
4.6115 | -99.1664
4.6152 | -99.1164
4.6188 | -99.0751
4.6224 | -99.0536
4.6261 | -99.0579
4.6297 | -99.0771
4.6333 | -99.1139
4.6369 | -99.1602
4.6406 | -99.2096
4.6442 | -99.2572
4.6478 | -99.2982
4.6515 | -99.337
4.6551 | -99.3678
4.6587 | -99.4053
4.6623 | -99.4429
4.666 | -99.4778
4.6696 | -99.5082
4.6732 | -99.5415
4.6769 | -99.5647
4.6805 | -99.5805
4.6841 | -99.6038
4.6877 | -99.6402
4.6914 | -99.6836
4.695 | -99.7255
4.6986 | -99.7757
4.7023 | -99.8303
4.7059 | -99.8886
4.7095 | -99.9186
4.7131 | -99.949
4.7168 | -99.9617
4.7204 | -99.9527
4.724 | -99.9327
4.7277 | -99.8993
4.7313 | -99.8559
4.7349 | -99.81
4.7385 | -99.768
4.7422 | -99.7345
4.7458 | -99.7146
4.7494 | -99.706
4.753 | -99.6989
4.7567 | -99.6887
4.7603 | -99.6816
4.7639 | -99.6789
4.7676 | -99.6818
4.7712 | -99.671
4.7748 | -99.651
4.7784 | -99.6303
4.7821 | -99.6065
4.7857 | -99.582
4.7893 | -99.5553
4.793 | -99.5353
4.7966 | -99.5125
4.8002 | -99.4836
4.8038 | -99.4555
4.8075 | -99.4299
4.8111 | -99.4069
4.8147 | -99.3902
4.8184 | -99.3849
4.822 | -99.3888
4.8256 | -99.4077
4.8292 | -99.4416
4.8329 | -99.4765
4.8365 | -99.5149
4.8401 | -99.5545
4.8438 | -99.5948
4.8474 | -99.6424
4.851 | -99.6909
4.8546 | -99.7413
4.8583 | -99.8014
4.8619 | -99.8601
4.8655 | -99.9203
4.8692 | -99.9772
4.8728 | -100.0298
4.8764 | -100.0732
4.88 | -100.1058
4.8837 | -100.1252
4.8873 | -100.1395
4.8909 | -100.1391
4.8946 | -100.1306
4.8982 | -100.1305
4.9018 | -100.1333
4.9054 | -100.1445
4.9091 | -100.1696
4.9127 | -100.2019
4.9163 | -100.2469
4.92 | -100.2906
4.9236 | -100.3356
4.9272 | -100.3818
4.9308 | -100.4191
4.9345 | -100.4435
4.9381 | -100.4668
4.9417 | -100.4723
4.9453 | -100.4624
4.949 | -100.4407
4.9526 | -100.4005
4.9562 | -100.3423
4.9599 | -100.2666
4.9635 | -100.1757
4.9671 | -100.0932
4.9707 | -100.0218
4.9744 | -99.9616
4.978 | -99.9245
4.9816 | -99.9062
4.9853 | -99.9101
4.9889 | -99.9247
4.9925 | -99.9366
4.9961 | -99.9547
4.9998 | -99.971
5.0034 | -99.9825
5.007 | -99.9891
5.0107 | -99.995
5.0143 | -100.0015
5.0179 | -100.0144
5.0215 | -100.0273
5.0252 | -100.0368
5.0288 | -100.0494
5.0324 | -100.0542
5.0361 | -100.0409
5.0397 | -100.0259
5.0433 | -100.0024
5.0469 | -99.9769
5.0506 | -99.9562
5.0542 | -99.943
5.0578 | -99.9426
5.0615 | -99.9529
5.0651 | -99.9784
5.0687 | -100.0216
5.0723 | -100.0721
5.076 | -100.1252
5.0796 | -100.1859
5.0832 | -100.2485
5.0869 | -100.2987
5.0905 | -100.3246
5.0941 | -100.3346
5.0977 | -100.3311
5.1014 | -100.317
5.105 | -100.2995
5.1086 | -100.2896
5.1123 | -100.2962
5.1159 | -100.3035
5.1195 | -100.3167
5.1231 | -100.3306
5.1268 | -100.3469
5.1304 | -100.3641
5.134 | -100.379
5.1376 | -100.3894
5.1413 | -100.405
5.1449 | -100.4215
5.1485 | -100.4356
5.1522 | -100.4587
5.1558 | -100.4857
5.1594 | -100.5248
5.163 | -100.5748
5.1667 | -100.6342
5.1703 | -100.696
5.1739 | -100.7543
5.1776 | -100.8016
5.1812 | -100.8385
5.1848 | -100.8614
5.1884 | -100.8605
5.1921 | -100.8485
5.1957 | -100.8279
5.1993 | -100.805
5.203 | -100.7784
5.2066 | -100.7503
5.2102 | -100.7174
5.2138 | -100.6821
5.2175 | -100.6476
5.2211 | -100.618
5.2247 | -100.594
5.2284 | -100.5729
5.232 | -100.5613
5.2356 | -100.5381
5.2392 | -100.5154
5.2429 | -100.489
5.2465 | -100.4575
5.2501 | -100.4157
5.2538 | -100.3734
5.2574 | -100.3319
5.261 | -100.2822
5.2646 | -100.2435
5.2683 | -100.2275
5.2719 | -100.2335
5.2755 | -100.2646
5.2792 | -100.3154
5.2828 | -100.3901
5.2864 | -100.4783
5.29 | -100.5832
5.2937 | -100.698
5.2973 | -100.8084
5.3009 | -100.9123
5.3045 | -101.0108
5.3082 | -101.0986
5.3118 | -101.1708
5.3154 | -101.2339
5.3191 | -101.2816
5.3227 | -101.3036
5.3263 | -101.3039
5.3299 | -101.2839
5.3336 | -101.2593
5.3372 | -101.2262
5.3408 | -101.1869
5.3445 | -101.1522
5.3481 | -101.1103
5.3517 | -101.0651
5.3553 | -101.0198
5.359 | -100.9797
5.3626 | -100.955
5.3662 | -100.9377
5.3699 | -100.9345
5.3735 | -100.9423
5.3771 | -100.9566
5.3807 | -100.9782
5.3844 | -101.0005
5.388 | -101.0223
5.3916 | -101.0391
5.3953 | -101.0528
5.3989 | -101.0613
5.4025 | -101.0595
5.4061 | -101.0497
5.4098 | -101.0333
5.4134 | -101.0125
5.417 | -100.9803
5.4207 | -100.9479
5.4243 | -100.9165
5.4279 | -100.8862
5.4315 | -100.8656
5.4352 | -100.8419
5.4388 | -100.8235
5.4424 | -100.8078
5.4461 | -100.782
5.4497 | -100.7581
5.4533 | -100.7401
5.4569 | -100.7215
5.4606 | -100.7124
5.4642 | -100.7174
5.4678 | -100.7327
5.4714 | -100.7621
5.4751 | -100.7981
5.4787 | -100.8464
5.4823 | -100.9028
5.486 | -100.9622
5.4896 | -101.028
5.4932 | -101.0936
5.4968 | -101.1513
5.5005 | -101.2063
5.5041 | -101.254
5.5077 | -101.3017
5.5114 | -101.3477
5.515 | -101.3768
5.5186 | -101.4001
5.5222 | -101.4064
5.5259 | -101.4078
5.5295 | -101.4028
5.5331 | -101.3765
5.5368 | -101.3318
5.5404 | -101.2767
5.544 | -101.208
5.5476 | -101.1267
5.5513 | -101.0494
5.5549 | -100.9787
5.5585 | -100.9151
5.5622 | -100.8592
5.5658 | -100.8118
5.5694 | -100.7725
5.573 | -100.7438
5.5767 | -100.7272
5.5803 | -100.7192
5.5839 | -100.7179
5.5876 | -100.7246
5.5912 | -100.7306
5.5948 | -100.7417
5.5984 | -100.761
5.6021 | -100.7799
5.6057 | -100.7978
5.6093 | -100.8069
5.613 | -100.8143
5.6166 | -100.8173
5.6202 | -100.813
5.6238 | -100.8091
5.6275 | -100.8053
5.6311 | -100.8038
5.6347 | -100.8063
5.6384 | -100.8104
5.642 | -100.8271
5.6456 | -100.8473
5.6492 | -100.8737
5.6529 | -100.9058
5.6565 | -100.9344
5.6601 | -100.9602
5.6637 | -100.9862
5.6674 | -101.015
5.671 | -101.044
5.6746 | -101.0697
5.6783 | -101.087
5.6819 | -101.0995
5.6855 | -101.1075
5.6891 | -101.1039
5.6928 | -101.0949
5.6964 | -101.0865
5.7 | -101.0838
5.7037 | -101.0859
5.7073 | -101.0874
5.7109 | -101.0922
5.7145 | -101.1028
5.7182 | -101.1147
5.7218 | -101.1293
5.7254 | -101.1422
5.7291 | -101.1505
5.7327 | -101.1603
5.7363 | -101.1722
5.7399 | -101.1877
5.7436 | -101.2126
5.7472 | -101.2492
5.7508 | -101.2934
5.7545 | -101.3687
5.7581 | -101.4223
5.7617 | -101.4882
5.7653 | -101.5563
5.769 | -101.6185
5.7726 | -101.6679
5.7762 | -101.7083
5.7799 | -101.7355
5.7835 | -101.7401
5.7871 | -101.722
5.7907 | -101.6873
5.7944 | -101.6413
5.798 | -101.5883
5.8016 | -101.5266
5.8053 | -101.4573
5.8089 | -101.3907
5.8125 | -101.3408
5.8161 | -101.3019
5.8198 | -101.2671
5.8234 | -101.2409
5.827 | -101.2214
5.8307 | -101.207
5.8343 | -101.1981
5.8379 | -101.1928
5.8415 | -101.1866
5.8452 | -101.1792
5.8488 | -101.1687
5.8524 | -101.1535
5.856 | -101.1394
5.8597 | -101.128
5.8633 | -101.1237
5.8669 | -101.123
5.8706 | -101.1205
5.8742 | -101.1254
5.8778 | -101.1436
5.8814 | -101.1754
5.8851 | -101.2112
5.8887 | -101.2467
5.8923 | -101.2913
5.896 | -101.3374
5.8996 | -101.3726
5.9032 | -101.3924
5.9068 | -101.404
5.9105 | -101.4049
5.9141 | -101.3985
5.9177 | -101.3927
5.9214 | -101.3852
5.925 | -101.3793
5.9286 | -101.3768
5.9322 | -101.3826
5.9359 | -101.3983
5.9395 | -101.4194
5.9431 | -101.4343
5.9468 | -101.4474
5.9504 | -101.4602
5.954 | -101.4649
5.9576 | -101.4638
5.9613 | -101.4652
5.9649 | -101.4601
5.9685 | -101.4466
5.9722 | -101.4322
5.9758 | -101.4259
5.9794 | -101.4304
5.983 | -101.4328
5.9867 | -101.4421
5.9903 | -101.457
5.9939 | -101.4756
5.9976 | -101.4977
6.0012 | -101.5165
6.0048 | -101.5337
6.0084 | -101.5406
6.0121 | -101.5368
6.0157 | -101.5252
6.0193 | -101.5077
6.0229 | -101.4801
6.0266 | -101.4462
6.0302 | -101.407
6.0338 | -101.3754
6.0375 | -101.347
6.0411 | -101.3205
6.0447 | -101.308
6.0483 | -101.3008
6.052 | -101.301
6.0556 | -101.3125
6.0592 | -101.3306
6.0629 | -101.3489
6.0665 | -101.3692
6.0701 | -101.3899
6.0737 | -101.4092
6.0774 | -101.4241
6.081 | -101.4406
6.0846 | -101.4606
6.0883 | -101.4796
6.0919 | -101.5053
6.0955 | -101.5306
6.0991 | -101.5621
6.1028 | -101.6013
6.1064 | -101.6416
6.11 | -101.681
6.1137 | -101.716
6.1173 | -101.7424
6.1209 | -101.76
6.1245 | -101.7681
6.1282 | -101.7632
6.1318 | -101.7523
6.1354 | -101.7266
6.1391 | -101.6966
6.1427 | -101.6598
6.1463 | -101.6194
6.1499 | -101.5816
6.1536 | -101.5473
6.1572 | -101.5182
6.1608 | -101.4991
6.1645 | -101.491
6.1681 | -101.4934
6.1717 | -101.4977
6.1753 | -101.503
6.179 | -101.5139
6.1826 | -101.5255
6.1862 | -101.5392
6.1898 | -101.5471
6.1935 | -101.5527
6.1971 | -101.5614
6.2007 | -101.5801
6.2044 | -101.6097
6.208 | -101.6503
6.2116 | -101.6958
6.2152 | -101.7426
6.2189 | -101.7953
6.2225 | -101.8402
6.2261 | -101.8742
6.2298 | -101.8885
6.2334 | -101.8819
6.237 | -101.8571
6.2406 | -101.8171
6.2443 | -101.7588
6.2479 | -101.6933
6.2515 | -101.6215
6.2552 | -101.5499
6.2588 | -101.4777
6.2624 | -101.415
6.266 | -101.359
6.2697 | -101.3116
6.2733 | -101.273
6.2769 | -101.242
6.2806 | -101.2417
6.2842 | -101.2283
6.2878 | -101.2211
6.2914 | -101.2197
6.2951 | -101.2258
6.2987 | -101.2408
6.3023 | -101.2595
6.306 | -101.2891
6.3096 | -101.325
6.3132 | -101.3624
6.3168 | -101.4046
6.3205 | -101.4543
6.3241 | -101.5123
6.3277 | -101.5752
6.3314 | -101.6402
6.335 | -101.7058
6.3386 | -101.7806
6.3422 | -101.8519
6.3459 | -101.9211
6.3495 | -101.9883
6.3531 | -102.0466
6.3568 | -102.099
6.3604 | -102.1472
6.364 | -102.1855
6.3676 | -102.2115
6.3713 | -102.2307
6.3749 | -102.2377
6.3785 | -102.2335
6.3821 | -102.2139
6.3858 | -102.1859
6.3894 | -102.1509
6.393 | -102.1098
6.3967 | -102.0684
6.4003 | -102.0257
6.4039 | -101.9809
6.4075 | -101.9396
6.4112 | -101.9063
6.4148 | -101.8784
6.4184 | -101.8491
6.4221 | -101.8214
6.4257 | -101.7974
6.4293 | -101.7755
6.4329 | -101.7606
6.4366 | -101.7524
6.4402 | -101.7556
6.4438 | -101.7714
6.4475 | -101.7905
6.4511 | -101.8197
6.4547 | -101.8588
6.4583 | -101.9081
6.462 | -101.9636
6.4656 | -102.0208
6.4692 | -102.0778
6.4729 | -102.134
6.4765 | -102.1825
6.4801 | -102.2292
6.4837 | -102.2667
6.4874 | -102.2904
6.491 | -102.3061
6.4946 | -102.3157
6.4983 | -102.3239
6.5019 | -102.3281
6.5055 | -102.329
6.5091 | -102.319
6.5128 | -102.3075
6.5164 | -102.3004
6.52 | -102.2934
6.5237 | -102.2898
6.5273 | -102.2897
6.5309 | -102.2878
6.5345 | -102.2866
6.5382 | -102.2849
6.5418 | -102.2838
6.5454 | -102.2933
6.5491 | -102.3031
6.5527 | -102.3097
6.5563 | -102.3232
6.5599 | -102.3449
6.5636 | -102.3764
6.5672 | -102.4122
6.5708 | -102.4559
6.5744 | -102.5092
6.5781 | -102.564
6.5817 | -102.6079
6.5853 | -102.6472
6.589 | -102.6789
6.5926 | -102.6944
6.5962 | -102.6907
6.5998 | -102.6733
6.6035 | -102.6452
6.6071 | -102.609
6.6107 | -102.5607
6.6144 | -102.5052
6.618 | -102.4482
6.6216 | -102.3908
6.6252 | -102.3398
6.6289 | -102.2938
6.6325 | -102.2573
6.6361 | -102.2272
6.6398 | -102.2031
6.6434 | -102.1844
6.647 | -102.1724
6.6506 | -102.171
6.6543 | -102.174
6.6579 | -102.187
6.6615 | -102.2059
6.6652 | -102.2218
6.6688 | -102.2348
6.6724 | -102.2455
6.676 | -102.2497
6.6797 | -102.256
6.6833 | -102.2526
6.6869 | -102.2342
6.6906 | -102.2118
6.6942 | -102.1885
6.6978 | -102.1623
6.7014 | -102.134
6.7051 | -102.1083
6.7087 | -102.0852
6.7123 | -102.0656
6.716 | -102.0477
6.7196 | -102.0323
6.7232 | -102.0209
6.7268 | -102.0209
6.7305 | -102.0259
6.7341 | -102.0378
6.7377 | -102.053
6.7413 | -102.0708
6.745 | -102.0888
6.7486 | -102.1
6.7522 | -102.1054
6.7559 | -102.1042
6.7595 | -102.0964
6.7631 | -102.0809
6.7667 | -102.0597
6.7704 | -102.0357
6.774 | -102.01
6.7776 | -101.9799
6.7813 | -101.9552
6.7849 | -101.9417
6.7885 | -101.9314
6.7921 | -101.9222
6.7958 | -101.916
6.7994 | -101.9196
6.803 | -101.948
6.8067 | -101.9632
6.8103 | -101.9853
6.8139 | -102.0104
6.8175 | -102.0426
6.8212 | -102.0773
6.8248 | -102.112
6.8284 | -102.1437
6.8321 | -102.1766
6.8357 | -102.2094
6.8393 | -102.2381
6.8429 | -102.2628
6.8466 | -102.2848
6.8502 | -102.2992
6.8538 | -102.3103
6.8575 | -102.317
6.8611 | -102.3198
6.8647 | -102.3215
6.8683 | -102.3269
6.872 | -102.3343
6.8756 | -102.3458
6.8792 | -102.3604
6.8829 | -102.3778
6.8865 | -102.3978
6.8901 | -102.4154
6.8937 | -102.4332
6.8974 | -102.4472
6.901 | -102.4546
6.9046 | -102.4646
6.9082 | -102.4715
6.9119 | -102.4773
6.9155 | -102.4844
6.9191 | -102.4919
6.9228 | -102.4986
6.9264 | -102.5119
6.93 | -102.5315
6.9336 | -102.5552
6.9373 | -102.5857
6.9409 | -102.6247
6.9445 | -102.6701
6.9482 | -102.716
6.9518 | -102.7662
6.9554 | -102.8199
6.959 | -102.8682
6.9627 | -102.9093
6.9663 | -102.943
6.9699 | -102.969
6.9736 | -102.9757
6.9772 | -102.968
6.9808 | -102.9469
6.9844 | -102.9156
6.9881 | -102.874
6.9917 | -102.8209
6.9953 | -102.7718
6.999 | -102.7229
7.0026 | -102.6807
7.0062 | -102.6444
7.0098 | -102.6117
7.0135 | -102.591
7.0171 | -102.5829
7.0207 | -102.5806
7.0244 | -102.594
7.028 | -102.6179
7.0316 | -102.6415
7.0352 | -102.6651
7.0389 | -102.6876
7.0425 | -102.7099
7.0461 | -102.7331
7.0498 | -102.7539
7.0534 | -102.7691
7.057 | -102.7757
7.0606 | -102.7763
7.0643 | -102.7773
7.0679 | -102.7716
7.0715 | -102.7549
7.0752 | -102.7364
7.0788 | -102.7228
7.0824 | -102.7106
7.086 | -102.7019
7.0897 | -102.6965
7.0933 | -102.6961
7.0969 | -102.6977
7.1005 | -102.6951
7.1042 | -102.6969
7.1078 | -102.6999
7.1114 | -102.6983
7.1151 | -102.6932
7.1187 | -102.6919
7.1223 | -102.6891
7.1259 | -102.6872
7.1296 | -102.6777
7.1332 | -102.6663
7.1368 | -102.6563
7.1405 | -102.6475
7.1441 | -102.6399
7.1477 | -102.6242
7.1513 | -102.6075
7.155 | -102.5937
7.1586 | -102.5859
7.1622 | -102.5819
7.1659 | -102.5822
7.1695 | -102.5894
7.1731 | -102.6053
7.1767 | -102.6304
7.1804 | -102.6585
7.184 | -102.6892
7.1876 | -102.7232
7.1913 | -102.7639
7.1949 | -102.8114
7.1985 | -102.86
7.2021 | -102.9148
7.2058 | -102.9699
7.2094 | -103.0178
7.213 | -103.0581
7.2167 | -103.0953
7.2203 | -103.1229
7.2239 | -103.1339
7.2275 | -103.1348
7.2312 | -103.1223
7.2348 | -103.1024
7.2384 | -103.0759
7.2421 | -103.0433
7.2457 | -103.003
7.2493 | -102.9592
7.2529 | -102.9202
7.2566 | -102.8818
7.2602 | -102.8429
7.2638 | -102.8046
7.2675 | -102.7738
7.2711 | -102.7526
7.2747 | -102.7359
7.2783 | -102.727
7.282 | -102.7248
7.2856 | -102.7346
7.2892 | -102.749
7.2928 | -102.7709
7.2965 | -102.7931
7.3001 | -102.8136
7.3037 | -102.8357
7.3074 | -102.8529
7.311 | -102.8702
7.3146 | -102.8855
7.3182 | -102.896
7.3219 | -102.9046
7.3255 | -102.9134
7.3291 | -102.9236
7.3328 | -102.9343
7.3364 | -102.948
7.34 | -102.9714
7.3436 | -102.9958
7.3473 | -103.0173
7.3509 | -103.0431
7.3545 | -103.0679
7.3582 | -103.0959
7.3618 | -103.1189
7.3654 | -103.1336
7.369 | -103.1488
7.3727 | -103.1607
7.3763 | -103.1626
7.3799 | -103.1584
7.3836 | -103.1495
7.3872 | -103.1436
7.3908 | -103.1397
7.3944 | -103.1321
7.3981 | -103.1277
7.4017 | -103.1192
7.4053 | -103.1107
7.409 | -103.1017
7.4126 | -103.0942
7.4162 | -103.086
7.4198 | -103.0804
7.4235 | -103.075
7.4271 | -103.0703
7.4307 | -103.0645
7.4344 | -103.0606
7.438 | -103.054
7.4416 | -103.0399
7.4452 | -103.0246
7.4489 | -103.0089
7.4525 | -102.9879
7.4561 | -102.9698
7.4597 | -102.9525
7.4634 | -102.9357
7.467 | -102.9243
7.4706 | -102.9094
7.4743 | -102.9031
7.4779 | -102.9023
7.4815 | -102.9034
7.4851 | -102.9088
7.4888 | -102.9146
7.4924 | -102.9178
7.496 | -102.922
7.4997 | -102.92
7.5033 | -102.9129
7.5069 | -102.9028
7.5105 | -102.8884
7.5142 | -102.8751
7.5178 | -102.8626
7.5214 | -102.8531
7.5251 | -102.8406
7.5287 | -102.8331
7.5323 | -102.8219
7.5359 | -102.8138
7.5396 | -102.8134
7.5432 | -102.819
7.5468 | -102.8247
7.5505 | -102.8348
7.5541 | -102.8498
7.5577 | -102.8656
7.5613 | -102.8794
7.565 | -102.8964
7.5686 | -102.9167
7.5722 | -102.9355
7.5759 | -102.9468
7.5795 | -102.956
7.5831 | -102.9696
7.5867 | -102.9782
7.5904 | -102.9895
7.594 | -102.9983
7.5976 | -103.0016
7.6013 | -103.0105
7.6049 | -103.0216
7.6085 | -103.0297
7.6121 | -103.0369
7.6158 | -103.0455
7.6194 | -103.0551
7.623 | -103.0698
7.6266 | -103.0823
7.6303 | -103.0951
7.6339 | -103.1058
7.6375 | -103.1232
7.6412 | -103.1397
7.6448 | -103.1595
7.6484 | -103.1748
7.652 | -103.1864
7.6557 | -103.1954
7.6593 | -103.203
7.6629 | -103.2132
7.6666 | -103.2221
7.6702 | -103.2289
7.6738 | -103.2372
7.6774 | -103.247
7.6811 | -103.2543
7.6847 | -103.268
7.6883 | -103.285
7.692 | -103.3014
7.6956 | -103.3168
7.6992 | -103.333
7.7028 | -103.3473
7.7065 | -103.3662
7.7101 | -103.389
7.7137 | -103.4048
7.7174 | -103.4173
7.721 | -103.4325
7.7246 | -103.4488
7.7282 | -103.4653
7.7319 | -103.4776
7.7355 | -103.4875
7.7391 | -103.5019
7.7428 | -103.5123
7.7464 | -103.5211
7.75 | -103.5258
7.7536 | -103.5285
7.7573 | -103.5305
7.7609 | -103.5313
7.7645 | -103.5291
7.7682 | -103.5244
7.7718 | -103.5242
7.7754 | -103.5218
7.779 | -103.5192
7.7827 | -103.5168
7.7863 | -103.5209
7.7899 | -103.5331
7.7936 | -103.5481
7.7972 | -103.5665
7.8008 | -103.5888
7.8044 | -103.6037
7.8081 | -103.6202
7.8117 | -103.6369
7.8153 | -103.6527
7.8189 | -103.6618
7.8226 | -103.6651
7.8262 | -103.6731
7.8298 | -103.678
7.8335 | -103.6807
7.8371 | -103.6839
7.8407 | -103.6831
7.8443 | -103.6769
7.848 | -103.6743
7.8516 | -103.6611
7.8552 | -103.6452
7.8589 | -103.6246
7.8625 | -103.6037
7.8661 | -103.5821
7.8697 | -103.5598
7.8734 | -103.5305
7.877 | -103.5015
7.8806 | -103.4763
7.8843 | -103.4509
7.8879 | -103.4238
7.8915 | -103.405
7.8951 | -103.3896
7.8988 | -103.3728
7.9024 | -103.3497
7.906 | -103.3287
7.9097 | -103.3134
7.9133 | -103.2925
7.9169 | -103.2694
7.9205 | -103.2507
7.9242 | -103.2255
7.9278 | -103.1965
7.9314 | -103.1665
7.9351 | -103.1398
7.9387 | -103.1166
7.9423 | -103.0941
7.9459 | -103.0763
7.9496 | -103.0634
7.9532 | -103.0555
7.9568 | -103.06
7.9605 | -103.0761
7.9641 | -103.0904
7.9677 | -103.1127
7.9713 | -103.1465
7.975 | -103.1895
7.9786 | -103.2368
7.9822 | -103.2811
7.9859 | -103.323
7.9895 | -103.3658
7.9931 | -103.4054
7.9967 | -103.4379
8.0004 | -103.4649
8.004 | -103.4894
8.0076 | -103.5106
8.0112 | -103.5228
8.0149 | -103.5364
8.0185 | -103.5503
8.0221 | -103.5602
8.0258 | -103.5677
8.0294 | -103.5744
8.033 | -103.5847
8.0366 | -103.5924
8.0403 | -103.6018
8.0439 | -103.6134
8.0475 | -103.6277
8.0512 | -103.6356
8.0548 | -103.6391
8.0584 | -103.6448
8.062 | -103.6479
8.0657 | -103.6497
8.0693 | -103.6478
8.0729 | -103.6412
8.0766 | -103.6263
8.0802 | -103.6041
8.0838 | -103.5818
8.0874 | -103.5527
8.0911 | -103.5164
8.0947 | -103.4772
8.0983 | -103.4375
8.102 | -103.4022
8.1056 | -103.3707
8.1092 | -103.3431
8.1128 | -103.3248
8.1165 | -103.315
8.1201 | -103.3104
8.1237 | -103.3112
8.1274 | -103.317
8.131 | -103.3329
8.1346 | -103.3529
8.1382 | -103.3801
8.1419 | -103.4131
8.1455 | -103.4425
8.1491 | -103.4772
8.1528 | -103.5177
8.1564 | -103.5566
8.16 | -103.5848
8.1636 | -103.6105
8.1673 | -103.6299
8.1709 | -103.6511
8.1745 | -103.665
8.1781 | -103.6691
8.1818 | -103.6698
8.1854 | -103.6608
8.189 | -103.6448
8.1927 | -103.629
8.1963 | -103.6103
8.1999 | -103.5927
8.2035 | -103.5709
8.2072 | -103.5555
8.2108 | -103.5393
8.2144 | -103.5271
8.2181 | -103.5189
8.2217 | -103.5079
8.2253 | -103.4988
8.2289 | -103.4966
8.2326 | -103.4971
8.2362 | -103.4993
8.2398 | -103.5041
8.2435 | -103.5174
8.2471 | -103.5339
8.2507 | -103.5462
8.2543 | -103.5569
8.258 | -103.5658
8.2616 | -103.5781
8.2652 | -103.5877
8.2689 | -103.591
8.2725 | -103.5886
8.2761 | -103.5897
8.2797 | -103.5937
8.2834 | -103.5886
8.287 | -103.5818
8.2906 | -103.5678
8.2943 | -103.5498
8.2979 | -103.5294
8.3015 | -103.5064
8.3051 | -103.4857
8.3088 | -103.4634
8.3124 | -103.4348
8.316 | -103.4093
8.3197 | -103.3829
8.3233 | -103.3585
8.3269 | -103.3284
8.3305 | -103.296
8.3342 | -103.2699
8.3378 | -103.2453
8.3414 | -103.2263
8.3451 | -103.2138
8.3487 | -103.2023
8.3523 | -103.1989
8.3559 | -103.2013
8.3596 | -103.2129
8.3632 | -103.2284
8.3668 | -103.2531
8.3704 | -103.287
8.3741 | -103.3452
8.3777 | -103.395
8.3813 | -103.4474
8.385 | -103.4993
8.3886 | -103.5529
8.3922 | -103.6117
8.3958 | -103.6716
8.3995 | -103.7299
8.4031 | -103.787
8.4067 | -103.8433
8.4104 | -103.8919
8.414 | -103.9342
8.4176 | -103.975
8.4212 | -104.0112
8.4249 | -104.0402
8.4285 | -104.063
8.4321 | -104.0767
8.4358 | -104.0861
8.4394 | -104.088
8.443 | -104.0853
8.4466 | -104.0786
8.4503 | -104.0719
8.4539 | -104.0674
8.4575 | -104.0568
8.4612 | -104.0476
8.4648 | -104.0327
8.4684 | -104.0142
8.472 | -103.996
8.4757 | -103.9734
8.4793 | -103.9452
8.4829 | -103.9128
8.4866 | -103.8835
8.4902 | -103.8568
8.4938 | -103.8307
8.4974 | -103.8043
8.5011 | -103.7782
8.5047 | -103.7518
8.5083 | -103.7343
8.512 | -103.7262
8.5156 | -103.717
8.5192 | -103.7065
8.5228 | -103.6998
8.5265 | -103.6946
8.5301 | -103.6867
8.5337 | -103.6806
8.5373 | -103.6806
8.541 | -103.6835
8.5446 | -103.6816
8.5482 | -103.6771
8.5519 | -103.6719
8.5555 | -103.6666
8.5591 | -103.6663
8.5627 | -103.6675
8.5664 | -103.6745
8.57 | -103.6821
8.5736 | -103.6902
8.5773 | -103.6993
8.5809 | -103.7097
8.5845 | -103.7172
8.5881 | -103.7212
8.5918 | -103.7296
8.5954 | -103.7347
8.599 | -103.7387
8.6027 | -103.7442
8.6063 | -103.7489
8.6099 | -103.751
8.6135 | -103.7454
8.6172 | -103.7345
8.6208 | -103.7243
8.6244 | -103.7168
8.6281 | -103.7095
8.6317 | -103.7041
8.6353 | -103.6951
8.6389 | -103.68
8.6426 | -103.6711
8.6462 | -103.6589
8.6498 | -103.6472
8.6535 | -103.6315
8.6571 | -103.6169
8.6607 | -103.6042
8.6643 | -103.5887
8.668 | -103.5676
8.6716 | -103.5519
8.6752 | -103.5388
8.6789 | -103.5261
8.6825 | -103.5155
8.6861 | -103.5006
8.6897 | -103.484
8.6934 | -103.4658
8.697 | -103.4477
8.7006 | -103.431
8.7043 | -103.4187
8.7079 | -103.4098
8.7115 | -103.403
8.7151 | -103.3988
8.7188 | -103.4002
8.7224 | -103.4027
8.726 | -103.4133
8.7296 | -103.4279
8.7333 | -103.4487
8.7369 | -103.4776
8.7405 | -103.5176
8.7442 | -103.5617
8.7478 | -103.6074
8.7514 | -103.6527
8.755 | -103.6979
8.7587 | -103.7456
8.7623 | -103.7921
8.7659 | -103.8333
8.7696 | -103.8656
8.7732 | -103.8918
8.7768 | -103.9099
8.7804 | -103.9298
8.7841 | -103.9412
8.7877 | -103.9409
8.7913 | -103.9374
8.795 | -103.9361
8.7986 | -103.9305
8.8022 | -103.9204
8.8058 | -103.9154
8.8095 | -103.9084
8.8131 | -103.9031
8.8167 | -103.8985
8.8204 | -103.8906
8.824 | -103.8882
8.8276 | -103.8905
8.8312 | -103.8883
8.8349 | -103.8836
8.8385 | -103.8785
8.8421 | -103.8771
8.8458 | -103.8708
8.8494 | -103.8625
8.853 | -103.8482
8.8566 | -103.8315
8.8603 | -103.8144
8.8639 | -103.789
8.8675 | -103.76
8.8712 | -103.7348
8.8748 | -103.7124
8.8784 | -103.6928
8.882 | -103.6746
8.8857 | -103.6594
8.8893 | -103.6491
8.8929 | -103.6466
8.8965 | -103.6677
8.9002 | -103.6813
8.9038 | -103.6961
8.9074 | -103.7129
8.9111 | -103.7329
8.9147 | -103.7565
8.9183 | -103.7835
8.9219 | -103.8077
8.9256 | -103.8315
8.9292 | -103.8549
8.9328 | -103.882
8.9365 | -103.9013
8.9401 | -103.9184
8.9437 | -103.9338
8.9473 | -103.9437
8.951 | -103.9533
8.9546 | -103.9568
8.9582 | -103.9534
8.9619 | -103.9522
8.9655 | -103.9531
8.9691 | -103.9593
8.9727 | -103.962
8.9764 | -103.9657
8.98 | -103.9752
8.9836 | -103.9891
8.9873 | -104.0055
8.9909 | -104.0192
8.9945 | -104.0409
8.9981 | -104.0642
9.0018 | -104.0819
9.0054 | -104.1001
9.009 | -104.1143
9.0127 | -104.1268
9.0163 | -104.1335
9.0199 | -104.1411
9.0235 | -104.146
9.0272 | -104.1431
9.0308 | -104.1314
9.0344 | -104.1207
9.0381 | -104.1121
9.0417 | -104.102
9.0453 | -104.0921
9.0489 | -104.0855
9.0526 | -104.0805
9.0562 | -104.0825
9.0598 | -104.0809
9.0634 | -104.079
9.0671 | -104.0801
9.0707 | -104.087
9.0743 | -104.093
9.078 | -104.1001
9.0816 | -104.1101
9.0852 | -104.1161
9.0888 | -104.1244
9.0925 | -104.1341
9.0961 | -104.1402
9.0997 | -104.1453
9.1034 | -104.1455
9.107 | -104.1464
9.1106 | -104.1497
9.1142 | -104.1521
9.1179 | -104.1539
9.1215 | -104.1579
9.1251 | -104.1595
9.1288 | -104.1551
9.1324 | -104.1536
9.136 | -104.1589
9.1396 | -104.1643
9.1433 | -104.1717
9.1469 | -104.1766
9.1505 | -104.1825
9.1542 | -104.1894
9.1578 | -104.1952
9.1614 | -104.2058
9.165 | -104.2181
9.1687 | -104.2269
9.1723 | -104.2351
9.1759 | -104.2464
9.1796 | -104.2551
9.1832 | -104.2635
9.1868 | -104.2733
9.1904 | -104.2846
9.1941 | -104.2995
9.1977 | -104.3136
9.2013 | -104.3285
9.205 | -104.3421
9.2086 | -104.3554
9.2122 | -104.3679
9.2158 | -104.3798
9.2195 | -104.3893
9.2231 | -104.396
9.2267 | -104.4022
9.2304 | -104.4052
9.234 | -104.3992
9.2376 | -104.3899
9.2412 | -104.384
9.2449 | -104.3712
9.2485 | -104.3478
9.2521 | -104.3247
9.2557 | -104.2998
9.2594 | -104.279
9.263 | -104.2564
9.2666 | -104.2308
9.2703 | -104.2043
9.2739 | -104.1779
9.2775 | -104.155
9.2811 | -104.1353
9.2848 | -104.1175
9.2884 | -104.0948
9.292 | -104.0766
9.2957 | -104.0572
9.2993 | -104.0379
9.3029 | -104.0193
9.3065 | -104.0017
9.3102 | -103.9845
9.3138 | -103.9702
9.3174 | -103.9588
9.3211 | -103.9514
9.3247 | -103.9464
9.3283 | -103.9466
9.3319 | -103.9508
9.3356 | -103.9604
9.3392 | -103.9759
9.3428 | -103.9946
9.3465 | -104.0139
9.3501 | -104.036
9.3537 | -104.0624
9.3573 | -104.0884
9.361 | -104.1129
9.3646 | -104.1408
9.3682 | -104.172
9.3719 | -104.1997
9.3755 | -104.2187
9.3791 | -104.237
9.3827 | -104.2588
9.3864 | -104.2792
9.39 | -104.2994
9.3936 | -104.3136
9.3973 | -104.3248
9.4009 | -104.3356
9.4045 | -104.347
9.4081 | -104.357
9.4118 | -104.3628
9.4154 | -104.3661
9.419 | -104.3611
9.4227 | -104.3568
9.4263 | -104.3525
9.4299 | -104.3459
9.4335 | -104.3371
9.4372 | -104.3234
9.4408 | -104.309
9.4444 | -104.2959
9.448 | -104.2794
9.4517 | -104.2619
9.4553 | -104.2458
9.4589 | -104.2295
9.4626 | -104.214
9.4662 | -104.2067
9.4698 | -104.2
9.4734 | -104.194
9.4771 | -104.1937
9.4807 | -104.201
9.4843 | -104.2115
9.488 | -104.2198
9.4916 | -104.2298
9.4952 | -104.2448
9.4988 | -104.2622
9.5025 | -104.2767
9.5061 | -104.2916
9.5097 | -104.3087
9.5134 | -104.3256
9.517 | -104.3482
9.5206 | -104.3703
9.5242 | -104.3857
9.5279 | -104.4008
9.5315 | -104.42
9.5351 | -104.4369
9.5388 | -104.451
9.5424 | -104.4615
9.546 | -104.4712
9.5496 | -104.4792
9.5533 | -104.4799
9.5569 | -104.4776
9.5605 | -104.4741
9.5642 | -104.4745
9.5678 | -104.4713
9.5714 | -104.4652
9.575 | -104.4597
9.5787 | -104.4589
9.5823 | -104.4555
9.5859 | -104.4512
9.5896 | -104.45
9.5932 | -104.4483
9.5968 | -104.4493
9.6004 | -104.4497
9.6041 | -104.4506
9.6077 | -104.4505
9.6113 | -104.4533
9.6149 | -104.4524
9.6186 | -104.4521
9.6222 | -104.4518
9.6258 | -104.4537
9.6295 | -104.4608
9.6331 | -104.4689
9.6367 | -104.4715
9.6403 | -104.4742
9.644 | -104.4817
9.6476 | -104.4892
9.6512 | -104.4962
9.6549 | -104.5026
9.6585 | -104.5007
9.6621 | -104.4999
9.6657 | -104.4999
9.6694 | -104.4972
9.673 | -104.4947
9.6766 | -104.4917
9.6803 | -104.4822
9.6839 | -104.4692
9.6875 | -104.4549
9.6911 | -104.4397
9.6948 | -104.4263
9.6984 | -104.4078
9.702 | -104.3843
9.7057 | -104.3603
9.7093 | -104.3315
9.7129 | -104.3003
9.7165 | -104.2712
9.7202 | -104.2378
9.7238 | -104.2047
9.7274 | -104.1707
9.7311 | -104.1376
9.7347 | -104.1063
9.7383 | -104.076
9.7419 | -104.0476
9.7456 | -104.022
9.7492 | -104.0006
9.7528 | -103.9784
9.7565 | -103.9688
9.7601 | -103.9636
9.7637 | -103.9559
9.7673 | -103.9495
9.771 | -103.9483
9.7746 | -103.952
9.7782 | -103.9609
9.7819 | -103.9726
9.7855 | -103.9875
9.7891 | -104.0027
9.7927 | -104.0182
9.7964 | -104.0336
9.8 | -104.0506
9.8036 | -104.0658
9.8072 | -104.0786
9.8109 | -104.0887
9.8145 | -104.0964
9.8181 | -104.1063
9.8218 | -104.1198
9.8254 | -104.1326
9.829 | -104.1446
9.8326 | -104.156
9.8363 | -104.1642
9.8399 | -104.1693
9.8435 | -104.1745
9.8472 | -104.1806
9.8508 | -104.192
9.8544 | -104.2017
9.858 | -104.2083
9.8617 | -104.2139
9.8653 | -104.2212
9.8689 | -104.2275
9.8726 | -104.2302
9.8762 | -104.2325
9.8798 | -104.23
9.8834 | -104.2281
9.8871 | -104.2247
9.8907 | -104.2193
9.8943 | -104.2184
9.898 | -104.2116
9.9016 | -104.204
9.9052 | -104.2028
9.9088 | -104.2027
9.9125 | -104.2056
9.9161 | -104.2098
9.9197 | -104.2182
9.9234 | -104.2292
9.927 | -104.2405
9.9306 | -104.2562
9.9342 | -104.2755
9.9379 | -104.3018
9.9415 | -104.3263
9.9451 | -104.3494
9.9488 | -104.3719
9.9524 | -104.3956
9.956 | -104.4139
9.9596 | -104.4275
9.9633 | -104.4372
9.9669 | -104.4424
9.9705 | -104.4443
9.9741 | -104.4414
9.9778 | -104.4362
9.9814 | -104.4308
9.985 | -104.422
9.9887 | -104.4072
9.9923 | -104.3905
9.9959 | -104.3703
9.9995 | -104.3491
10.0032 | -104.3252
10.0068 | -104.3015
10.0104 | -104.2764
10.0141 | -104.2547
10.0177 | -104.2357
10.0213 | -104.2142
10.0249 | -104.1909
10.0286 | -104.1669
10.0322 | -104.1483
10.0358 | -104.1318
10.0395 | -104.1156
10.0431 | -104.0981
10.0467 | -104.0831
10.0503 | -104.0715
10.054 | -104.0585
10.0576 | -104.0426
10.0612 | -104.0296
10.0649 | -104.0187
10.0685 | -104.0098
10.0721 | -104.0039
10.0757 | -103.9971
10.0794 | -103.9899
10.083 | -103.985
10.0866 | -103.9857
10.0903 | -103.991
10.0939 | -103.9994
10.0975 | -104.0098
10.1011 | -104.0206
10.1048 | -104.0375
10.1084 | -104.0588
10.112 | -104.0792
10.1157 | -104.1007
10.1193 | -104.126
10.1229 | -104.1539
10.1265 | -104.1826
10.1302 | -104.2111
10.1338 | -104.2424
10.1374 | -104.2749
10.141 | -104.3055
10.1447 | -104.3384
10.1483 | -104.3738
10.1519 | -104.4112
10.1556 | -104.4489
10.1592 | -104.4835
10.1628 | -104.5108
10.1664 | -104.5412
10.1701 | -104.5747
10.1737 | -104.6111
10.1773 | -104.6463
10.181 | -104.6728
10.1846 | -104.6934
10.1882 | -104.7151
10.1918 | -104.7346
10.1955 | -104.75
10.1991 | -104.7604
10.2027 | -104.7701
10.2064 | -104.7776
10.21 | -104.7847
10.2136 | -104.7884
10.2172 | -104.7888
10.2209 | -104.7866
10.2245 | -104.7885
10.2281 | -104.79
10.2318 | -104.7882
10.2354 | -104.7858
10.239 | -104.78
10.2426 | -104.7756
10.2463 | -104.7742
10.2499 | -104.7789
10.2535 | -104.7826
10.2572 | -104.7813
10.2608 | -104.7843
10.2644 | -104.7919
10.268 | -104.7992
10.2717 | -104.8049
10.2753 | -104.8121
10.2789 | -104.8171
10.2826 | -104.8186
10.2862 | -104.8228
10.2898 | -104.8281
10.2934 | -104.8363
10.2971 | -104.8411
10.3007 | -104.8396
10.3043 | -104.8368
10.3079 | -104.8358
10.3116 | -104.8342
10.3152 | -104.8248
10.3188 | -104.8147
10.3225 | -104.7996
10.3261 | -104.7853
10.3297 | -104.7744
10.3333 | -104.7661
10.337 | -104.7557
10.3406 | -104.7477
10.3442 | -104.7399
10.3479 | -104.733
10.3515 | -104.7284
10.3551 | -104.7216
10.3587 | -104.7152
10.3624 | -104.7111
10.366 | -104.7083
10.3696 | -104.7081
10.3733 | -104.7061
10.3769 | -104.7059
10.3805 | -104.7037
10.3841 | -104.703
10.3878 | -104.7033
10.3914 | -104.7016
10.395 | -104.6985
10.3987 | -104.6948
10.4023 | -104.6945
10.4059 | -104.6946
10.4095 | -104.693
10.4132 | -104.6899
10.4168 | -104.6888
10.4204 | -104.6914
10.4241 | -104.6939
10.4277 | -104.6995
10.4313 | -104.7075
10.4349 | -104.7082
10.4386 | -104.709
10.4422 | -104.7057
10.4458 | -104.7057
10.4495 | -104.71
10.4531 | -104.7141
10.4567 | -104.7175
10.4603 | -104.7202
10.464 | -104.7188
10.4676 | -104.7109
10.4712 | -104.7095
10.4749 | -104.7069
10.4785 | -104.703
10.4821 | -104.7001
10.4857 | -104.6972
10.4894 | -104.6949
10.493 | -104.6914
10.4966 | -104.6893
10.5002 | -104.688
10.5039 | -104.6842
10.5075 | -104.6845
10.5111 | -104.683
10.5148 | -104.6827
10.5184 | -104.6833
10.522 | -104.6853
10.5256 | -104.6883
10.5293 | -104.6927
10.5329 | -104.7
10.5365 | -104.7092
10.5402 | -104.7173
10.5438 | -104.7225
10.5474 | -104.7307
10.551 | -104.7384
10.5547 | -104.7468
10.5583 | -104.7583
10.5619 | -104.7705
10.5656 | -104.782
10.5692 | -104.7886
10.5728 | -104.794
10.5764 | -104.8005
10.5801 | -104.8059
10.5837 | -104.812
10.5873 | -104.8221
10.591 | -104.8303
10.5946 | -104.8412
10.5982 | -104.856
10.6018 | -104.8703
10.6055 | -104.8827
10.6091 | -104.8914
10.6127 | -104.9062
10.6164 | -104.9272
10.62 | -104.9473
10.6236 | -104.9605
10.6272 | -104.9744
10.6309 | -104.9924
10.6345 | -105.0077
10.6381 | -105.0151
10.6418 | -105.0193
10.6454 | -105.0246
10.649 | -105.0318
10.6526 | -105.0348
10.6563 | -105.0364
10.6599 | -105.038
10.6635 | -105.033
10.6671 | -105.0283
10.6708 | -105.0231
10.6744 | -105.0153
10.678 | -105.0017
10.6817 | -104.986
10.6853 | -104.9732
10.6889 | -104.961
10.6925 | -104.9468
10.6962 | -104.9366
10.6998 | -104.9333
10.7034 | -104.9268
10.7071 | -104.9195
10.7107 | -104.914
10.7143 | -104.9098
10.7179 | -104.9097
10.7216 | -104.9132
10.7252 | -104.9185
10.7288 | -104.9222
10.7325 | -104.9265
10.7361 | -104.9317
10.7397 | -104.9371
10.7433 | -104.9448
10.747 | -104.9527
10.7506 | -104.959
10.7542 | -104.9604
10.7579 | -104.9646
10.7615 | -104.9719
10.7651 | -104.9821
10.7687 | -104.9866
10.7724 | -104.9904
10.776 | -104.9926
10.7796 | -104.9963
10.7833 | -104.9953
10.7869 | -104.9934
10.7905 | -104.9909
10.7941 | -104.9877
10.7978 | -104.9833
10.8014 | -104.9728
10.805 | -104.9611
10.8087 | -104.9493
10.8123 | -104.9363
10.8159 | -104.9193
10.8195 | -104.9063
10.8232 | -104.891
10.8268 | -104.8736
10.8304 | -104.8579
10.8341 | -104.8477
10.8377 | -104.8395
10.8413 | -104.8269
10.8449 | -104.8138
10.8486 | -104.8022
10.8522 | -104.7889
10.8558 | -104.7756
10.8594 | -104.767
10.8631 | -104.7574
10.8667 | -104.7447
10.8703 | -104.7317
10.874 | -104.72
10.8776 | -104.7067
10.8812 | -104.6971
10.8848 | -104.6896
10.8885 | -104.678
10.8921 | -104.6667
10.8957 | -104.6576
10.8994 | -104.6456
10.903 | -104.6321
10.9066 | -104.6231
10.9102 | -104.6153
10.9139 | -104.6089
10.9175 | -104.5996
10.9211 | -104.5899
10.9248 | -104.5782
10.9284 | -104.5648
10.932 | -104.5575
10.9356 | -104.5482
10.9393 | -104.5381
10.9429 | -104.528
10.9465 | -104.5215
10.9502 | -104.517
10.9538 | -104.5096
10.9574 | -104.503
10.961 | -104.4982
10.9647 | -104.4983
10.9683 | -104.4988
10.9719 | -104.5037
10.9756 | -104.5104
10.9792 | -104.5237
10.9828 | -104.5426
10.9864 | -104.5655
10.9901 | -104.6029
10.9937 | -104.6259
10.9973 | -104.6527
11.001 | -104.6806
11.0046 | -104.7122
11.0082 | -104.744
11.0118 | -104.7788
11.0155 | -104.8108
11.0191 | -104.8397
11.0227 | -104.8729
11.0264 | -104.9064
11.03 | -104.9361
11.0336 | -104.9569
11.0372 | -104.9763
11.0409 | -104.9941
11.0445 | -105.0082
11.0481 | -105.0216
11.0517 | -105.0309
11.0554 | -105.0349
11.059 | -105.0424
11.0626 | -105.0504
11.0663 | -105.0558
11.0699 | -105.0623
11.0735 | -105.0648
11.0771 | -105.0688
11.0808 | -105.0711
11.0844 | -105.0794
11.088 | -105.0857
11.0917 | -105.0908
11.0953 | -105.0968
11.0989 | -105.1017
11.1025 | -105.1069
11.1062 | -105.1085
11.1098 | -105.1114
11.1134 | -105.1144
11.1171 | -105.1181
11.1207 | -105.1202
11.1243 | -105.1201
11.1279 | -105.1191
11.1316 | -105.1166
11.1352 | -105.1095
11.1388 | -105.0964
11.1425 | -105.0812
11.1461 | -105.0698
11.1497 | -105.056
11.1533 | -105.0424
11.157 | -105.0317
11.1606 | -105.0166
11.1642 | -105.0023
11.1679 | -104.9854
11.1715 | -104.9692
11.1751 | -104.9501
11.1787 | -104.9337
11.1824 | -104.92
11.186 | -104.9034
11.1896 | -104.8894
11.1933 | -104.8794
11.1969 | -104.8748
11.2005 | -104.874
11.2041 | -104.871
11.2078 | -104.8676
11.2114 | -104.8666
11.215 | -104.8672
11.2186 | -104.8666
11.2223 | -104.8638
11.2259 | -104.8647
11.2295 | -104.8668
11.2332 | -104.8702
11.2368 | -104.876
11.2404 | -104.8867
11.244 | -104.8981
11.2477 | -104.907
11.2513 | -104.9176
11.2549 | -104.9275
11.2586 | -104.9417
11.2622 | -104.9612
11.2658 | -104.9831
11.2694 | -105.007
11.2731 | -105.0363
11.2767 | -105.0672
11.2803 | -105.0959
11.284 | -105.1214
11.2876 | -105.149
11.2912 | -105.1778
11.2948 | -105.2082
11.2985 | -105.2383
11.3021 | -105.2669
11.3057 | -105.2916
11.3094 | -105.304
11.313 | -105.3165
11.3166 | -105.3277
11.3202 | -105.3322
11.3239 | -105.3337
11.3275 | -105.3351
11.3311 | -105.3327
11.3348 | -105.3256
11.3384 | -105.3205
11.342 | -105.3119
11.3456 | -105.2983
11.3493 | -105.2878
11.3529 | -105.277
11.3565 | -105.2628
11.3602 | -105.246
11.3638 | -105.2263
11.3674 | -105.2071
11.371 | -105.1903
11.3747 | -105.1753
11.3783 | -105.1607
11.3819 | -105.1505
11.3856 | -105.1375
11.3892 | -105.1202
11.3928 | -105.1067
11.3964 | -105.0937
11.4001 | -105.0815
11.4037 | -105.0666
11.4073 | -105.0487
11.4109 | -105.0285
11.4146 | -105.0096
11.4182 | -104.9888
11.4218 | -104.9663
11.4255 | -104.9399
11.4291 | -104.9106
11.4327 | -104.8806
11.4363 | -104.8465
11.44 | -104.8126
11.4436 | -104.7791
11.4472 | -104.7486
11.4509 | -104.7151
11.4545 | -104.6805
11.4581 | -104.6446
11.4617 | -104.6109
11.4654 | -104.581
11.469 | -104.5559
11.4726 | -104.5371
11.4763 | -104.5193
11.4799 | -104.5039
11.4835 | -104.496
11.4871 | -104.4958
11.4908 | -104.4967
11.4944 | -104.5019
11.498 | -104.5109
11.5017 | -104.5221
11.5053 | -104.5358
11.5089 | -104.5554
11.5125 | -104.5828
11.5162 | -104.6045
11.5198 | -104.6263
11.5234 | -104.6512
11.5271 | -104.679
11.5307 | -104.707
11.5343 | -104.7376
11.5379 | -104.7688
11.5416 | -104.7977
11.5452 | -104.828
11.5488 | -104.8543
11.5525 | -104.8855
11.5561 | -104.9192
11.5597 | -104.9464
11.5633 | -104.97
11.567 | -104.9931
11.5706 | -105.02
11.5742 | -105.0509
11.5778 | -105.0813
11.5815 | -105.1081
11.5851 | -105.1323
11.5887 | -105.1514
11.5924 | -105.1704
11.596 | -105.1876
11.5996 | -105.2027
11.6032 | -105.2196
11.6069 | -105.2346
11.6105 | -105.2461
11.6141 | -105.2531
11.6178 | -105.2602
11.6214 | -105.2652
11.625 | -105.2648
11.6286 | -105.2593
11.6323 | -105.2518
11.6359 | -105.2391
11.6395 | -105.2231
11.6432 | -105.2073
11.6468 | -105.1884
11.6504 | -105.1678
11.654 | -105.1423
11.6577 | -105.1193
11.6613 | -105.0957
11.6649 | -105.0724
11.6686 | -105.0456
11.6722 | -105.022
11.6758 | -105.0004
11.6794 | -104.98
11.6831 | -104.9608
11.6867 | -104.9419
11.6903 | -104.9274
11.694 | -104.9146
11.6976 | -104.9057
11.7012 | -104.8998
11.7048 | -104.8967
11.7085 | -104.8963
11.7121 | -104.8989
11.7157 | -104.9064
11.7194 | -104.9183
11.723 | -104.9291
11.7266 | -104.9423
11.7302 | -104.9592
11.7339 | -104.9807
11.7375 | -105.0043
11.7411 | -105.027
11.7447 | -105.0522
11.7484 | -105.0765
11.752 | -105.1001
11.7556 | -105.1267
11.7593 | -105.1564
11.7629 | -105.1833
11.7665 | -105.2079
11.7701 | -105.2319
11.7738 | -105.2483
11.7774 | -105.2658
11.781 | -105.284
11.7847 | -105.2993
11.7883 | -105.3114
11.7919 | -105.3229
11.7955 | -105.3322
11.7992 | -105.3389
11.8028 | -105.3386
11.8064 | -105.342
11.8101 | -105.3418
11.8137 | -105.3344
11.8173 | -105.329
11.8209 | -105.3231
11.8246 | -105.3165
11.8282 | -105.3073
11.8318 | -105.3
11.8355 | -105.2942
11.8391 | -105.2888
11.8427 | -105.2816
11.8463 | -105.2735
11.85 | -105.2661
11.8536 | -105.2579
11.8572 | -105.2509
11.8609 | -105.246
11.8645 | -105.2444
11.8681 | -105.2428
11.8717 | -105.2426
11.8754 | -105.2427
11.879 | -105.2459
11.8826 | -105.248
11.8863 | -105.2493
11.8899 | -105.2496
11.8935 | -105.2565
11.8971 | -105.2604
11.9008 | -105.2643
11.9044 | -105.2646
11.908 | -105.2637
11.9117 | -105.2604
11.9153 | -105.2539
11.9189 | -105.2465
11.9225 | -105.2399
11.9262 | -105.2279
11.9298 | -105.214
11.9334 | -105.198
11.937 | -105.1779
11.9407 | -105.1576
11.9443 | -105.1393
11.9479 | -105.1162
11.9516 | -105.0909
11.9552 | -105.0713
11.9588 | -105.0502
11.9624 | -105.025
11.9661 | -105.0058
11.9697 | -104.9895
11.9733 | -104.9736
11.977 | -104.9529
11.9806 | -104.9367
11.9842 | -104.9231
11.9878 | -104.9074
11.9915 | -104.894
11.9951 | -104.882
11.9987 | -104.8712
12.0024 | -104.8619
12.006 | -104.8519
12.0096 | -104.843
12.0132 | -104.8361
12.0169 | -104.8313
12.0205 | -104.8311
12.0241 | -104.8309
12.0278 | -104.8285
12.0314 | -104.8224
12.035 | -104.8145
12.0386 | -104.8085
12.0423 | -104.8005
12.0459 | -104.7959
12.0495 | -104.7928
12.0532 | -104.7873
12.0568 | -104.7833
12.0604 | -104.7807
12.064 | -104.7799
12.0677 | -104.7732
12.0713 | -104.7707
12.0749 | -104.7689
12.0786 | -104.767
12.0822 | -104.7602
12.0858 | -104.7547
12.0894 | -104.7536
12.0931 | -104.7559
12.0967 | -104.7573
12.1003 | -104.7616
12.1039 | -104.7661
12.1076 | -104.7708
12.1112 | -104.7784
12.1148 | -104.7832
12.1185 | -104.7865
12.1221 | -104.7908
12.1257 | -104.7922
12.1293 | -104.7911
12.133 | -104.7916
12.1366 | -104.7935
12.1402 | -104.791
12.1439 | -104.7863
12.1475 | -104.7858
12.1511 | -104.7873
12.1547 | -104.786
12.1584 | -104.7881
12.162 | -104.7923
12.1656 | -104.7954
12.1693 | -104.8028
12.1729 | -104.8137
12.1765 | -104.8252
12.1801 | -104.8398
12.1838 | -104.8558
12.1874 | -104.8725
12.191 | -104.8899
12.1947 | -104.9064
12.1983 | -104.9248
12.2019 | -104.9436
12.2055 | -104.9672
12.2092 | -104.9918
12.2128 | -105.0105
12.2164 | -105.0276
12.2201 | -105.0465
12.2237 | -105.0633
12.2273 | -105.0771
12.2309 | -105.09
12.2346 | -105.0987
12.2382 | -105.1038
12.2418 | -105.1085
12.2455 | -105.1141
12.2491 | -105.1182
12.2527 | -105.1224
12.2563 | -105.1254
12.26 | -105.1295
12.2636 | -105.1278
12.2672 | -105.1214
12.2709 | -105.1123
12.2745 | -105.1007
12.2781 | -105.0849
12.2817 | -105.0679
12.2854 | -105.0556
12.289 | -105.0396
12.2926 | -105.0239
12.2962 | -105.0096
12.2999 | -104.9965
12.3035 | -104.983
12.3071 | -104.9678
12.3108 | -104.957
12.3144 | -104.951
12.318 | -104.9476
12.3216 | -104.9467
12.3253 | -104.9459
12.3289 | -104.9458
12.3325 | -104.9494
12.3362 | -104.9554
12.3398 | -104.9622
12.3434 | -104.9711
12.347 | -104.979
12.3507 | -104.9862
12.3543 | -104.9964
12.3579 | -105.0078
12.3616 | -105.0208
12.3652 | -105.0346
12.3688 | -105.0461
12.3724 | -105.0604
12.3761 | -105.0772
12.3797 | -105.0934
12.3833 | -105.1066
12.387 | -105.1212
12.3906 | -105.1368
12.3942 | -105.1498
12.3978 | -105.1654
12.4015 | -105.1803
12.4051 | -105.1948
12.4087 | -105.2081
12.4124 | -105.2213
12.416 | -105.2332
12.4196 | -105.2423
12.4232 | -105.2455
12.4269 | -105.2469
12.4305 | -105.2467
12.4341 | -105.2472
12.4378 | -105.2467
12.4414 | -105.2457
12.445 | -105.2424
12.4486 | -105.2354
12.4523 | -105.2324
12.4559 | -105.229
12.4595 | -105.2239
12.4632 | -105.2156
12.4668 | -105.2083
12.4704 | -105.2017
12.474 | -105.1949
12.4777 | -105.1886
12.4813 | -105.1859
12.4849 | -105.1842
12.4885 | -105.1802
12.4922 | -105.1757
12.4958 | -105.1723
12.4994 | -105.1711
12.5031 | -105.1693
12.5067 | -105.1714
12.5103 | -105.1712
12.5139 | -105.168
12.5176 | -105.1651
12.5212 | -105.1655
12.5248 | -105.1686
12.5285 | -105.169
12.5321 | -105.1643
12.5357 | -105.1606
12.5393 | -105.159
12.543 | -105.1594
12.5466 | -105.1634
12.5502 | -105.1658
12.5539 | -105.1658
12.5575 | -105.1658
12.5611 | -105.1668
12.5647 | -105.1708
12.5684 | -105.1762
12.572 | -105.184
12.5756 | -105.1941
12.5793 | -105.201
12.5829 | -105.2075
12.5865 | -105.2139
12.5901 | -105.2237
12.5938 | -105.2365
12.5974 | -105.2453
12.601 | -105.2537
12.6047 | -105.2661
12.6083 | -105.2771
12.6119 | -105.2841
12.6155 | -105.2897
12.6192 | -105.2959
12.6228 | -105.3014
12.6264 | -105.3076
12.6301 | -105.314
12.6337 | -105.3157
12.6373 | -105.3148
12.6409 | -105.3128
12.6446 | -105.3095
12.6482 | -105.3097
12.6518 | -105.3085
12.6554 | -105.3054
12.6591 | -105.2979
12.6627 | -105.2869
12.6663 | -105.2778
12.67 | -105.2691
12.6736 | -105.2596
12.6772 | -105.252
12.6808 | -105.2451
12.6845 | -105.2429
12.6881 | -105.2399
12.6917 | -105.2424
12.6954 | -105.2471
12.699 | -105.2538
12.7026 | -105.2605
12.7062 | -105.2691
12.7099 | -105.2841
12.7135 | -105.3047
12.7171 | -105.3246
12.7208 | -105.3447
12.7244 | -105.3642
12.728 | -105.3854
12.7316 | -105.4078
12.7353 | -105.4332
12.7389 | -105.4581
12.7425 | -105.4799
12.7462 | -105.4945
12.7498 | -105.5085
12.7534 | -105.5182
12.757 | -105.5243
12.7607 | -105.5258
12.7643 | -105.5229
12.7679 | -105.5163
12.7716 | -105.5094
12.7752 | -105.4969
12.7788 | -105.4799
12.7824 | -105.4631
12.7861 | -105.4444
12.7897 | -105.4222
12.7933 | -105.3993
12.797 | -105.3772
12.8006 | -105.3534
12.8042 | -105.3306
12.8078 | -105.3111
12.8115 | -105.2892
12.8151 | -105.2684
12.8187 | -105.2488
12.8224 | -105.2331
12.826 | -105.2166
12.8296 | -105.2012
12.8332 | -105.1841
12.8369 | -105.166
12.8405 | -105.1487
12.8441 | -105.1377
12.8477 | -105.1325
12.8514 | -105.1225
12.855 | -105.1102
12.8586 | -105.1008
12.8623 | -105.0927
12.8659 | -105.0876
12.8695 | -105.0831
12.8731 | -105.0819
12.8768 | -105.0813
12.8804 | -105.0812
12.884 | -105.0823
12.8877 | -105.0852
12.8913 | -105.09
12.8949 | -105.0962
12.8985 | -105.1017
12.9022 | -105.1081
12.9058 | -105.1161
12.9094 | -105.1251
12.9131 | -105.1368
12.9167 | -105.1486
12.9203 | -105.1601
12.9239 | -105.172
12.9276 | -105.1869
12.9312 | -105.2031
12.9348 | -105.2164
12.9385 | -105.2293
12.9421 | -105.2419
12.9457 | -105.256
12.9493 | -105.2715
12.953 | -105.2825
12.9566 | -105.29
12.9602 | -105.2959
12.9639 | -105.3018
12.9675 | -105.3056
12.9711 | -105.308
12.9747 | -105.3084
12.9784 | -105.3068
12.982 | -105.3062
12.9856 | -105.3039
12.9893 | -105.2976
12.9929 | -105.2877
12.9965 | -105.2827
13.0001 | -105.2777
13.0038 | -105.2693
13.0074 | -105.2617
13.011 | -105.2594
13.0146 | -105.2577
13.0183 | -105.256
13.0219 | -105.2594
13.0255 | -105.2636
13.0292 | -105.265
13.0328 | -105.268
13.0364 | -105.2736
13.04 | -105.2841
13.0437 | -105.2937
13.0473 | -105.3033
13.0509 | -105.3117
13.0546 | -105.3232
13.0582 | -105.3372
13.0618 | -105.3527
13.0654 | -105.3679
13.0691 | -105.3848
13.0727 | -105.403
13.0763 | -105.4223
13.08 | -105.4404
13.0836 | -105.4525
13.0872 | -105.4676
13.0908 | -105.4852
13.0945 | -105.502
13.0981 | -105.518
13.1017 | -105.5292
13.1054 | -105.5382
13.109 | -105.5486
13.1126 | -105.5614
13.1162 | -105.571
13.1199 | -105.5819
13.1235 | -105.5924
13.1271 | -105.6002
13.1308 | -105.606
13.1344 | -105.6109
13.138 | -105.6112
13.1416 | -105.6116
13.1453 | -105.6121
13.1489 | -105.6097
13.1525 | -105.6046
13.1562 | -105.5999
13.1598 | -105.5936
13.1634 | -105.586
13.167 | -105.576
13.1707 | -105.5615
13.1743 | -105.5471
13.1779 | -105.5329
13.1815 | -105.5172
13.1852 | -105.4974
13.1888 | -105.4774
13.1924 | -105.4637
13.1961 | -105.4476
13.1997 | -105.4282
13.2033 | -105.4056
13.2069 | -105.3872
13.2106 | -105.3723
13.2142 | -105.3595
13.2178 | -105.3487
13.2215 | -105.3356
13.2251 | -105.3262
13.2287 | -105.3162
13.2323 | -105.3085
13.236 | -105.3031
13.2396 | -105.2996
13.2432 | -105.2951
13.2469 | -105.2921
13.2505 | -105.2895
13.2541 | -105.2904
13.2577 | -105.2894
13.2614 | -105.2894
13.265 | -105.289
13.2686 | -105.2902
13.2723 | -105.2922
13.2759 | -105.2967
13.2795 | -105.3016
13.2831 | -105.3049
13.2868 | -105.306
13.2904 | -105.3087
13.294 | -105.3137
13.2977 | -105.3205
13.3013 | -105.3252
13.3049 | -105.3273
13.3085 | -105.3316
13.3122 | -105.3321
13.3158 | -105.3337
13.3194 | -105.3394
13.3231 | -105.3472
13.3267 | -105.3535
13.3303 | -105.3613
13.3339 | -105.3679
13.3376 | -105.3702
13.3412 | -105.3744
13.3448 | -105.3777
13.3485 | -105.3794
13.3521 | -105.3788
13.3557 | -105.3794
13.3593 | -105.3817
13.363 | -105.3842
13.3666 | -105.3859
13.3702 | -105.3858
13.3738 | -105.3843
13.3775 | -105.3838
13.3811 | -105.3839
13.3847 | -105.3795
13.3884 | -105.3767
13.392 | -105.3729
13.3956 | -105.3724
13.3992 | -105.3692
13.4029 | -105.3635
13.4065 | -105.3582
13.4101 | -105.3523
13.4138 | -105.3503
13.4174 | -105.3469
13.421 | -105.3454
13.4246 | -105.3446
13.4283 | -105.3431
13.4319 | -105.342
13.4355 | -105.3407
13.4392 | -105.3419
13.4428 | -105.3445
13.4464 | -105.3455
13.45 | -105.3489
13.4537 | -105.3541
13.4573 | -105.3622
13.4609 | -105.3696
13.4646 | -105.3797
13.4682 | -105.3917
13.4718 | -105.4026
13.4754 | -105.4138
13.4791 | -105.4252
13.4827 | -105.4341
13.4863 | -105.4463
13.49 | -105.4618
13.4936 | -105.4796
13.4972 | -105.4999
13.5008 | -105.5183
13.5045 | -105.5383
13.5081 | -105.5587
13.5117 | -105.5777
13.5154 | -105.5936
13.519 | -105.6088
13.5226 | -105.6238
13.5262 | -105.6387
13.5299 | -105.6521
13.5335 | -105.6594
13.5371 | -105.6662
13.5407 | -105.6718
13.5444 | -105.678
13.548 | -105.679
13.5516 | -105.6759
13.5553 | -105.6719
13.5589 | -105.6648
13.5625 | -105.6556
13.5661 | -105.6445
13.5698 | -105.629
13.5734 | -105.6095
13.577 | -105.5916
13.5807 | -105.5734
13.5843 | -105.5562
13.5879 | -105.5391
13.5915 | -105.523
13.5952 | -105.505
13.5988 | -105.4869
13.6024 | -105.4698
13.6061 | -105.4551
13.6097 | -105.4361
13.6133 | -105.4181
13.6169 | -105.3977
13.6206 | -105.3791
13.6242 | -105.3625
13.6278 | -105.3476
13.6315 | -105.333
13.6351 | -105.3155
13.6387 | -105.3019
13.6423 | -105.2891
13.646 | -105.2758
13.6496 | -105.2599
13.6532 | -105.2446
13.6569 | -105.2311
13.6605 | -105.2194
13.6641 | -105.208
13.6677 | -105.1951
13.6714 | -105.1843
13.675 | -105.176
13.6786 | -105.1695
13.6823 | -105.1578
13.6859 | -105.1438
13.6895 | -105.1318
13.6931 | -105.1203
13.6968 | -105.1108
13.7004 | -105.1035
13.704 | -105.0958
13.7077 | -105.0907
13.7113 | -105.0891
13.7149 | -105.0882
13.7185 | -105.0862
13.7222 | -105.0844
13.7258 | -105.0855
13.7294 | -105.0883
13.733 | -105.0889
13.7367 | -105.0904
13.7403 | -105.0948
13.7439 | -105.1027
13.7476 | -105.1129
13.7512 | -105.1223
13.7548 | -105.1326
13.7584 | -105.1457
13.7621 | -105.1597
13.7657 | -105.1744
13.7693 | -105.1943
13.773 | -105.2121
13.7766 | -105.2267
13.7802 | -105.2437
13.7838 | -105.2605
13.7875 | -105.2771
13.7911 | -105.2946
13.7947 | -105.3146
13.7984 | -105.3333
13.802 | -105.3509
13.8056 | -105.3697
13.8092 | -105.3884
13.8129 | -105.4089
13.8165 | -105.4288
13.8201 | -105.4447
13.8238 | -105.4558
13.8274 | -105.4668
13.831 | -105.4812
13.8346 | -105.4946
13.8383 | -105.5045
13.8419 | -105.5142
13.8455 | -105.5209
13.8492 | -105.5287
13.8528 | -105.5365
13.8564 | -105.5449
13.86 | -105.5513
13.8637 | -105.5568
13.8673 | -105.5641
13.8709 | -105.5698
13.8746 | -105.5754
13.8782 | -105.5792
13.8818 | -105.5811
13.8854 | -105.5837
13.8891 | -105.5855
13.8927 | -105.5839
13.8963 | -105.584
13.9 | -105.5803
13.9036 | -105.5782
13.9072 | -105.5752
13.9108 | -105.5714
13.9145 | -105.567
13.9181 | -105.5618
13.9217 | -105.5601
13.9253 | -105.5593
13.929 | -105.5573
13.9326 | -105.5577
13.9362 | -105.5561
13.9399 | -105.5509
13.9435 | -105.5454
13.9471 | -105.542
13.9507 | -105.5407
13.9544 | -105.5395
13.958 | -105.5326
13.9616 | -105.5046
13.9653 | -105.4584
13.9689 | -105.4315
13.9725 | -105.4083
13.9761 | -105.3855
13.9798 | -105.3567
13.9834 | -105.3223
13.987 | -105.2857
13.9907 | -105.2471
13.9943 | -105.2031
13.9979 | -105.1573
14.0015 | -105.1102
14.0052 | -105.0596
14.0088 | -105.0072
14.0124 | -104.9524
14.0161 | -104.8961
14.0197 | -104.8402
14.0233 | -104.7873
14.0269 | -104.7338
14.0306 | -104.6795
14.0342 | -104.6271
14.0378 | -104.5792
14.0415 | -104.5376
14.0451 | -104.5032
14.0487 | -104.4731
14.0523 | -104.4461
14.056 | -104.424
14.0596 | -104.4081
14.0632 | -104.3997
14.0669 | -104.3986
14.0705 | -104.4056
14.0741 | -104.4185
14.0777 | -104.4389
14.0814 | -104.4665
14.085 | -104.5002
14.0886 | -104.538
14.0922 | -104.5824
14.0959 | -104.6309
14.0995 | -104.6843
14.1031 | -104.7355
14.1068 | -104.7879
14.1104 | -104.8375
14.114 | -104.8904
14.1176 | -104.9443
14.1213 | -104.9941
14.1249 | -105.0406
14.1285 | -105.0867
14.1322 | -105.1164
14.1358 | -105.154
14.1394 | -105.1912
14.143 | -105.2239
14.1467 | -105.2561
14.1503 | -105.2836
14.1539 | -105.3085
14.1576 | -105.3319
14.1612 | -105.3588
14.1648 | -105.4009
14.1684 | -105.4484
14.1721 | -105.4649
14.1757 | -105.475
14.1793 | -105.4823
14.183 | -105.4914
14.1866 | -105.5027
14.1902 | -105.5129
14.1938 | -105.5208
14.1975 | -105.5276
14.2011 | -105.53
14.2047 | -105.5328
14.2084 | -105.5356
14.212 | -105.5364
14.2156 | -105.5353
14.2192 | -105.5374
14.2229 | -105.5395
14.2265 | -105.5399
14.2301 | -105.5401
14.2338 | -105.5388
14.2374 | -105.54
14.241 | -105.5419
14.2446 | -105.5393
14.2483 | -105.5358
14.2519 | -105.5346
14.2555 | -105.5373
14.2591 | -105.5416
14.2628 | -105.5463
14.2664 | -105.5503
14.27 | -105.5533
14.2737 | -105.5538
14.2773 | -105.5573
14.2809 | -105.5617
14.2845 | -105.5672
14.2882 | -105.5724
14.2918 | -105.5743
14.2954 | -105.5786
14.2991 | -105.5841
14.3027 | -105.5928
14.3063 | -105.605
14.3099 | -105.616
14.3136 | -105.6288
14.3172 | -105.6416
14.3208 | -105.6509
14.3245 | -105.6607
14.3281 | -105.6722
14.3317 | -105.6857
14.3353 | -105.6989
14.339 | -105.7125
14.3426 | -105.7265
14.3462 | -105.7386
14.3499 | -105.7501
14.3535 | -105.7558
14.3571 | -105.7604
14.3607 | -105.7645
14.3644 | -105.7651
14.368 | -105.76
14.3716 | -105.7533
14.3753 | -105.7444
14.3789 | -105.7309
14.3825 | -105.7161
14.3861 | -105.702
14.3898 | -105.6856
14.3934 | -105.665
14.397 | -105.6425
14.4007 | -105.6199
14.4043 | -105.5961
14.4079 | -105.5711
14.4115 | -105.5406
14.4152 | -105.5055
14.4188 | -105.4755
14.4224 | -105.4448
14.4261 | -105.4104
14.4297 | -105.3747
14.4333 | -105.3412
14.4369 | -105.3085
14.4406 | -105.2798
14.4442 | -105.2513
14.4478 | -105.2222
14.4514 | -105.1986
14.4551 | -105.1757
14.4587 | -105.1547
14.4623 | -105.1373
14.466 | -105.1218
14.4696 | -105.1094
14.4732 | -105.101
14.4768 | -105.0947
14.4805 | -105.0946
14.4841 | -105.0963
14.4877 | -105.1
14.4914 | -105.108
14.495 | -105.1165
14.4986 | -105.1308
14.5022 | -105.1445
14.5059 | -105.1562
14.5095 | -105.1704
14.5131 | -105.1864
14.5168 | -105.2055
14.5204 | -105.2276
14.524 | -105.2446
14.5276 | -105.2667
14.5313 | -105.2904
14.5349 | -105.3101
14.5385 | -105.3272
14.5422 | -105.3461
14.5458 | -105.366
14.5494 | -105.3843
14.553 | -105.4029
14.5567 | -105.4224
14.5603 | -105.4389
14.5639 | -105.4543
14.5676 | -105.467
14.5712 | -105.4792
14.5748 | -105.4903
14.5784 | -105.4995
14.5821 | -105.5088
14.5857 | -105.517
14.5893 | -105.5243
14.593 | -105.5275
14.5966 | -105.5294
14.6002 | -105.5333
14.6038 | -105.5359
14.6075 | -105.5365
14.6111 | -105.5342
14.6147 | -105.5349
14.6183 | -105.536
14.622 | -105.5387
14.6256 | -105.5377
14.6292 | -105.5378
14.6329 | -105.5387
14.6365 | -105.5407
14.6401 | -105.54
14.6437 | -105.5375
14.6474 | -105.5335
14.651 | -105.5299
14.6546 | -105.5227
14.6583 | -105.5199
14.6619 | -105.5178
14.6655 | -105.5167
14.6691 | -105.5144
14.6728 | -105.5101
14.6764 | -105.5078
14.68 | -105.5058
14.6837 | -105.5038
14.6873 | -105.5021
14.6909 | -105.4992
14.6945 | -105.4954
14.6982 | -105.4911
14.7018 | -105.4869
14.7054 | -105.4871
14.7091 | -105.4897
14.7127 | -105.4956
14.7163 | -105.5051
14.7199 | -105.5127
14.7236 | -105.5204
14.7272 | -105.5289
14.7308 | -105.5381
14.7345 | -105.5508
14.7381 | -105.5632
14.7417 | -105.5725
14.7453 | -105.5845
14.749 | -105.5982
14.7526 | -105.6094
14.7562 | -105.6197
14.7599 | -105.6299
14.7635 | -105.642
14.7671 | -105.653
14.7707 | -105.6647
14.7744 | -105.6746
14.778 | -105.6821
14.7816 | -105.6853
14.7853 | -105.6888
14.7889 | -105.6922
14.7925 | -105.6913
14.7961 | -105.6888
14.7998 | -105.6858
14.8034 | -105.6828
14.807 | -105.6781
14.8106 | -105.672
14.8143 | -105.6653
14.8179 | -105.6567
14.8215 | -105.6501
14.8252 | -105.6402
14.8288 | -105.6313
14.8324 | -105.6208
14.836 | -105.6093
14.8397 | -105.5972
14.8433 | -105.5883
14.8469 | -105.5787
14.8506 | -105.5706
14.8542 | -105.5651
14.8578 | -105.5599
14.8614 | -105.5568
14.8651 | -105.5514
14.8687 | -105.5447
14.8723 | -105.5401
14.876 | -105.537
14.8796 | -105.5361
14.8832 | -105.5366
14.8868 | -105.5379
14.8905 | -105.539
14.8941 | -105.542
14.8977 | -105.5466
14.9014 | -105.5521
14.905 | -105.5578
14.9086 | -105.5643
14.9122 | -105.5726
14.9159 | -105.5797
14.9195 | -105.5875
14.9231 | -105.5949
14.9268 | -105.5994
14.9304 | -105.6023
14.934 | -105.6043
14.9376 | -105.6094
14.9413 | -105.6116
14.9449 | -105.6152
14.9485 | -105.6183
14.9522 | -105.6211
14.9558 | -105.6243
14.9594 | -105.6283
14.963 | -105.6277
14.9667 | -105.6287
14.9703 | -105.6311
14.9739 | -105.6329
14.9776 | -105.632
14.9812 | -105.6296
14.9848 | -105.6251
14.9884 | -105.6155
14.9921 | -105.6088
14.9957 | -105.6059
14.9993 | -105.599
15.0029 | -105.5908
15.0066 | -105.5847
15.0102 | -105.5783
15.0138 | -105.5675
15.0175 | -105.5577
15.0211 | -105.5474
15.0247 | -105.5347
15.0283 | -105.5199
15.032 | -105.5084
15.0356 | -105.5001
15.0392 | -105.4909
15.0429 | -105.4779
15.0465 | -105.4642
15.0501 | -105.4533
15.0537 | -105.4434
15.0574 | -105.4296
15.061 | -105.4183
15.0646 | -105.4074
15.0683 | -105.3989
15.0719 | -105.3926
15.0755 | -105.3858
15.0791 | -105.3847
15.0828 | -105.3831
15.0864 | -105.3786
15.09 | -105.3717
15.0937 | -105.3656
15.0973 | -105.3581
15.1009 | -105.3535
15.1045 | -105.3515
15.1082 | -105.3484
15.1118 | -105.3457
15.1154 | -105.3449
15.1191 | -105.343
15.1227 | -105.3425
15.1263 | -105.3422
15.1299 | -105.3415
15.1336 | -105.3376
15.1372 | -105.3336
15.1408 | -105.3351
15.1445 | -105.3379
15.1481 | -105.3434
15.1517 | -105.345
15.1553 | -105.3482
15.159 | -105.3495
15.1626 | -105.352
15.1662 | -105.3562
15.1698 | -105.3617
15.1735 | -105.3661
15.1771 | -105.3749
15.1807 | -105.3833
15.1844 | -105.3919
15.188 | -105.401
15.1916 | -105.4094
15.1952 | -105.419
15.1989 | -105.4283
15.2025 | -105.4401
15.2061 | -105.4576
15.2098 | -105.4704
15.2134 | -105.4806
15.217 | -105.4916
15.2206 | -105.5026
15.2243 | -105.5104
15.2279 | -105.517
15.2315 | -105.5252
15.2352 | -105.5296
15.2388 | -105.5318
15.2424 | -105.5367
15.246 | -105.5415
15.2497 | -105.542
15.2533 | -105.5358
15.2569 | -105.5309
15.2606 | -105.5278
15.2642 | -105.5255
15.2678 | -105.5203
15.2714 | -105.5142
15.2751 | -105.5091
15.2787 | -105.5006
15.2823 | -105.4918
15.286 | -105.4823
15.2896 | -105.4729
15.2932 | -105.462
15.2968 | -105.4509
15.3005 | -105.4445
15.3041 | -105.4404
15.3077 | -105.4386
15.3114 | -105.4381
15.315 | -105.4411
15.3186 | -105.4417
15.3222 | -105.4408
15.3259 | -105.4411
15.3295 | -105.4398
15.3331 | -105.4437
15.3368 | -105.4476
15.3404 | -105.4518
15.344 | -105.4584
15.3476 | -105.4656
15.3513 | -105.4758
15.3549 | -105.4859
15.3585 | -105.4938
15.3621 | -105.5024
15.3658 | -105.509
15.3694 | -105.5175
15.373 | -105.5266
15.3767 | -105.5344
15.3803 | -105.5396
15.3839 | -105.5422
15.3875 | -105.5467
15.3912 | -105.5518
15.3948 | -105.5582
15.3984 | -105.5588
15.4021 | -105.5585
15.4057 | -105.5564
15.4093 | -105.5503
15.4129 | -105.5441
15.4166 | -105.5375
15.4202 | -105.5285
15.4238 | -105.5212
15.4275 | -105.5113
15.4311 | -105.5002
15.4347 | -105.4886
15.4383 | -105.4788
15.442 | -105.4677
15.4456 | -105.46
15.4492 | -105.4516
15.4529 | -105.443
15.4565 | -105.4369
15.4601 | -105.4294
15.4637 | -105.4196
15.4674 | -105.4106
15.471 | -105.4031
15.4746 | -105.3966
15.4783 | -105.3868
15.4819 | -105.3773
15.4855 | -105.3716
15.4891 | -105.367
15.4928 | -105.3617
15.4964 | -105.3574
15.5 | -105.3577
15.5037 | -105.356
15.5073 | -105.3528
15.5109 | -105.351
15.5145 | -105.3518
15.5182 | -105.3504
15.5218 | -105.3513
15.5254 | -105.3513
15.529 | -105.3498
15.5327 | -105.3496
15.5363 | -105.3507
15.5399 | -105.3517
15.5436 | -105.3549
15.5472 | -105.36
15.5508 | -105.3648
15.5544 | -105.3644
15.5581 | -105.3676
15.5617 | -105.3722
15.5653 | -105.3748
15.569 | -105.3803
15.5726 | -105.3838
15.5762 | -105.3886
15.5798 | -105.3923
15.5835 | -105.3964
15.5871 | -105.4024
15.5907 | -105.4076
15.5944 | -105.4121
15.598 | -105.4169
15.6016 | -105.4218
15.6052 | -105.4266
15.6089 | -105.4311
15.6125 | -105.4358
15.6161 | -105.4391
15.6198 | -105.4458
15.6234 | -105.4532
15.627 | -105.4612
15.6306 | -105.4697
15.6343 | -105.4744
15.6379 | -105.4799
15.6415 | -105.4885
15.6452 | -105.4898
15.6488 | -105.4914
15.6524 | -105.4952
15.656 | -105.4976
15.6597 | -105.4962
15.6633 | -105.4924
15.6669 | -105.4864
15.6706 | -105.4804
15.6742 | -105.4713
15.6778 | -105.4611
15.6814 | -105.4518
15.6851 | -105.4414
15.6887 | -105.4297
15.6923 | -105.4146
15.6959 | -105.4032
15.6996 | -105.3923
15.7032 | -105.3736
15.7068 | -105.3614
15.7105 | -105.3454
15.7141 | -105.33
15.7177 | -105.316
15.7213 | -105.3022
15.725 | -105.2874
15.7286 | -105.2764
15.7322 | -105.2662
15.7359 | -105.2553
15.7395 | -105.2448
15.7431 | -105.2373
15.7467 | -105.226
15.7504 | -105.2177
15.754 | -105.212
15.7576 | -105.2056
15.7613 | -105.1992
15.7649 | -105.1961
15.7685 | -105.1945
15.7721 | -105.1912
15.7758 | -105.1901
15.7794 | -105.191
15.783 | -105.1929
15.7867 | -105.1939
15.7903 | -105.1938
15.7939 | -105.1945
15.7975 | -105.198
15.8012 | -105.1995
15.8048 | -105.2037
15.8084 | -105.209
15.8121 | -105.2125
15.8157 | -105.2188
15.8193 | -105.2251
15.8229 | -105.2332
15.8266 | -105.2421
15.8302 | -105.2494
15.8338 | -105.2549
15.8375 | -105.2611
15.8411 | -105.2664
15.8447 | -105.2721
15.8483 | -105.2771
15.852 | -105.2815
15.8556 | -105.2863
15.8592 | -105.2953
15.8629 | -105.2998
15.8665 | -105.3023
15.8701 | -105.3073
15.8737 | -105.312
15.8774 | -105.3165
15.881 | -105.3214
15.8846 | -105.3285
15.8882 | -105.334
15.8919 | -105.3401
15.8955 | -105.3478
15.8991 | -105.3534
15.9028 | -105.3603
15.9064 | -105.3658
15.91 | -105.3712
15.9136 | -105.3742
15.9173 | -105.3787
15.9209 | -105.3805
15.9245 | -105.3826
15.9282 | -105.3861
15.9318 | -105.3896
15.9354 | -105.3899
15.939 | -105.3903
15.9427 | -105.3909
15.9463 | -105.3915
15.9499 | -105.3906
15.9536 | -105.3887
15.9572 | -105.3836
15.9608 | -105.3783
15.9644 | -105.3745
15.9681 | -105.3715
15.9717 | -105.3702
15.9753 | -105.369
15.979 | -105.3675
15.9826 | -105.3658
15.9862 | -105.3648
15.9898 | -105.3644
15.9935 | -105.3642
15.9971 | -105.3621
16.0007 | -105.3633
16.0044 | -105.3691
16.008 | -105.3731
16.0116 | -105.3756
16.0152 | -105.3797
16.0189 | -105.3856
16.0225 | -105.3905
16.0261 | -105.3952
16.0298 | -105.4035
16.0334 | -105.4127
16.037 | -105.42
16.0406 | -105.4281
16.0443 | -105.4355
16.0479 | -105.4451
16.0515 | -105.4528
16.0552 | -105.4597
16.0588 | -105.469
16.0624 | -105.4785
16.066 | -105.4852
16.0697 | -105.491
16.0733 | -105.4955
16.0769 | -105.4994
16.0805 | -105.5056
16.0842 | -105.5128
16.0878 | -105.5176
16.0914 | -105.5237
16.0951 | -105.5295
16.0987 | -105.534
16.1023 | -105.5364
16.1059 | -105.536
16.1096 | -105.5357
16.1132 | -105.5355
16.1168 | -105.5391
16.1205 | -105.5374
16.1241 | -105.5343
16.1277 | -105.5352
16.1313 | -105.5328
16.135 | -105.5311
16.1386 | -105.5289
16.1422 | -105.5269
16.1459 | -105.5243
16.1495 | -105.5197
16.1531 | -105.5161
16.1567 | -105.5123
16.1604 | -105.5077
16.164 | -105.5073
16.1676 | -105.5072
16.1713 | -105.5059
16.1749 | -105.5054
16.1785 | -105.5072
16.1821 | -105.5099
16.1858 | -105.5156
16.1894 | -105.5215
16.193 | -105.5244
16.1967 | -105.53
16.2003 | -105.5363
16.2039 | -105.5436
16.2075 | -105.5558
16.2112 | -105.5687
16.2148 | -105.5802
16.2184 | -105.5929
16.2221 | -105.6086
16.2257 | -105.6296
16.2293 | -105.6479
16.2329 | -105.6667
16.2366 | -105.6823
16.2402 | -105.701
16.2438 | -105.7205
16.2474 | -105.7419
16.2511 | -105.7644
16.2547 | -105.785
16.2583 | -105.8063
16.262 | -105.8263
16.2656 | -105.8445
16.2692 | -105.8647
16.2728 | -105.8851
16.2765 | -105.9046
16.2801 | -105.9215
16.2837 | -105.9391
16.2874 | -105.9604
16.291 | -105.9808
16.2946 | -106.0006
16.2982 | -106.0203
16.3019 | -106.0387
16.3055 | -106.0568
16.3091 | -106.0747
16.3128 | -106.0868
16.3164 | -106.0975
16.32 | -106.1101
16.3236 | -106.1231
16.3273 | -106.1354
16.3309 | -106.1464
16.3345 | -106.1536
16.3382 | -106.1603
16.3418 | -106.1659
16.3454 | -106.1693
16.349 | -106.1699
16.3527 | -106.1724
16.3563 | -106.1713
16.3599 | -106.1623
16.3636 | -106.1541
16.3672 | -106.1458
16.3708 | -106.1352
16.3744 | -106.1227
16.3781 | -106.1077
16.3817 | -106.0927
16.3853 | -106.0759
16.389 | -106.056
16.3926 | -106.0343
16.3962 | -106.0113
16.3998 | -105.9882
16.4035 | -105.9671
16.4071 | -105.9468
16.4107 | -105.9242
16.4143 | -105.9017
16.418 | -105.8796
16.4216 | -105.8583
16.4252 | -105.8399
16.4289 | -105.8201
16.4325 | -105.799
16.4361 | -105.7821
16.4397 | -105.7612
16.4434 | -105.7436
16.447 | -105.7322
16.4506 | -105.7219
16.4543 | -105.7115
16.4579 | -105.7027
16.4615 | -105.6952
16.4651 | -105.6895
16.4688 | -105.6842
16.4724 | -105.6814
16.476 | -105.6791
16.4797 | -105.6769
16.4833 | -105.6774
16.4869 | -105.6803
16.4905 | -105.6856
16.4942 | -105.6908
16.4978 | -105.6944
16.5014 | -105.6987
16.5051 | -105.7033
16.5087 | -105.7101
16.5123 | -105.7164
16.5159 | -105.7241
16.5196 | -105.7328
16.5232 | -105.7386
16.5268 | -105.7447
16.5305 | -105.7528
16.5341 | -105.763
16.5377 | -105.7754
16.5413 | -105.7842
16.545 | -105.7927
16.5486 | -105.802
16.5522 | -105.8099
16.5559 | -105.8195
16.5595 | -105.8273
16.5631 | -105.8361
16.5667 | -105.8449
16.5704 | -105.8554
16.574 | -105.8655
16.5776 | -105.8739
16.5813 | -105.881
16.5849 | -105.8884
16.5885 | -105.8951
16.5921 | -105.901
16.5958 | -105.9112
16.5994 | -105.9189
16.603 | -105.9266
16.6066 | -105.9331
16.6103 | -105.9359
16.6139 | -105.9372
16.6175 | -105.9403
16.6212 | -105.9426
16.6248 | -105.9454
16.6284 | -105.949
16.632 | -105.9503
16.6357 | -105.9517
16.6393 | -105.9497
16.6429 | -105.9489
16.6466 | -105.9528
16.6502 | -105.9564
16.6538 | -105.9591
16.6574 | -105.9633
16.6611 | -105.9678
16.6647 | -105.9688
16.6683 | -105.9681
16.672 | -105.969
16.6756 | -105.975
16.6792 | -105.9799
16.6828 | -105.9819
16.6865 | -105.9842
16.6901 | -105.9885
16.6937 | -105.992
16.6974 | -105.996
16.701 | -105.9998
16.7046 | -106.0045
16.7082 | -106.0096
16.7119 | -106.0171
16.7155 | -106.0254
16.7191 | -106.0281
16.7228 | -106.0302
16.7264 | -106.0358
16.73 | -106.0419
16.7336 | -106.0502
16.7373 | -106.0592
16.7409 | -106.0691
16.7445 | -106.0755
16.7482 | -106.0811
16.7518 | -106.0878
16.7554 | -106.0926
16.759 | -106.0983
16.7627 | -106.1066
16.7663 | -106.1142
16.7699 | -106.1195
16.7736 | -106.1209
16.7772 | -106.122
16.7808 | -106.1252
16.7844 | -106.1262
16.7881 | -106.1263
16.7917 | -106.1271
16.7953 | -106.124
16.7989 | -106.1201
16.8026 | -106.1138
16.8062 | -106.1049
16.8098 | -106.0951
16.8135 | -106.0847
16.8171 | -106.0729
16.8207 | -106.0609
16.8243 | -106.0495
16.828 | -106.0362
16.8316 | -106.0211
16.8352 | -106.0019
16.8389 | -105.9827
16.8425 | -105.9635
16.8461 | -105.9448
16.8497 | -105.9285
16.8534 | -105.9136
16.857 | -105.8962
16.8606 | -105.8777
16.8643 | -105.8595
16.8679 | -105.8424
16.8715 | -105.8234
16.8751 | -105.8076
16.8788 | -105.7965
16.8824 | -105.786
16.886 | -105.7753
16.8897 | -105.7644
16.8933 | -105.7576
16.8969 | -105.7516
16.9005 | -105.7468
16.9042 | -105.7498
16.9078 | -105.757
16.9114 | -105.7633
16.9151 | -105.7658
16.9187 | -105.7703
16.9223 | -105.7776
16.9259 | -105.7872
16.9296 | -105.7966
16.9332 | -105.8084
16.9368 | -105.8215
16.9405 | -105.8334
16.9441 | -105.8434
16.9477 | -105.8516
16.9513 | -105.8599
16.955 | -105.8712
16.9586 | -105.8841
16.9622 | -105.8949
16.9658 | -105.9018
16.9695 | -105.9082
16.9731 | -105.9152
16.9767 | -105.9213
16.9804 | -105.9259
16.984 | -105.9325
16.9876 | -105.9368
16.9912 | -105.9392
16.9949 | -105.9395
16.9985 | -105.9394
17.0021 | -105.9377
17.0058 | -105.9345
17.0094 | -105.9289
17.013 | -105.9235
17.0166 | -105.9212
17.0203 | -105.9178
17.0239 | -105.9117
17.0275 | -105.9058
17.0312 | -105.9004
17.0348 | -105.8942
17.0384 | -105.8856
17.042 | -105.8747
17.0457 | -105.8633
17.0493 | -105.8545
17.0529 | -105.845
17.0566 | -105.8387
17.0602 | -105.8308
17.0638 | -105.8229
17.0674 | -105.8112
17.0711 | -105.8027
17.0747 | -105.7978
17.0783 | -105.7934
17.082 | -105.7863
17.0856 | -105.7794
17.0892 | -105.7762
17.0928 | -105.7746
17.0965 | -105.7697
17.1001 | -105.7653
17.1037 | -105.7642
17.1074 | -105.7631
17.111 | -105.7629
17.1146 | -105.7648
17.1182 | -105.7633
17.1219 | -105.761
17.1255 | -105.758
17.1291 | -105.7568
17.1328 | -105.7587
17.1364 | -105.7567
17.14 | -105.7563
17.1436 | -105.7572
17.1473 | -105.7574
17.1509 | -105.7559
17.1545 | -105.7548
17.1581 | -105.7577
17.1618 | -105.7625
17.1654 | -105.7664
17.169 | -105.7674
17.1727 | -105.77
17.1763 | -105.7723
17.1799 | -105.7757
17.1835 | -105.7795
17.1872 | -105.7857
17.1908 | -105.7929
17.1944 | -105.7997
17.1981 | -105.804
17.2017 | -105.8089
17.2053 | -105.8138
17.2089 | -105.8212
17.2126 | -105.829
17.2162 | -105.8354
17.2198 | -105.8418
17.2235 | -105.847
17.2271 | -105.8522
17.2307 | -105.8606
17.2343 | -105.8705
17.238 | -105.8806
17.2416 | -105.8899
17.2452 | -105.8987
17.2489 | -105.9056
17.2525 | -105.9173
17.2561 | -105.9294
17.2597 | -105.94
17.2634 | -105.9511
17.267 | -105.9622
17.2706 | -105.9739
17.2743 | -105.9846
17.2779 | -105.9936
17.2815 | -106.0038
17.2851 | -106.0149
17.2888 | -106.0293
17.2924 | -106.0428
17.296 | -106.0535
17.2997 | -106.0632
17.3033 | -106.0688
17.3069 | -106.0766
17.3105 | -106.0826
17.3142 | -106.0923
17.3178 | -106.1005
17.3214 | -106.1063
17.325 | -106.1128
17.3287 | -106.1212
17.3323 | -106.1279
17.3359 | -106.1308
17.3396 | -106.1338
17.3432 | -106.1396
17.3468 | -106.1428
17.3504 | -106.1454
17.3541 | -106.1482
17.3577 | -106.1483
17.3613 | -106.1489
17.365 | -106.152
17.3686 | -106.1535
17.3722 | -106.1529
17.3758 | -106.1508
17.3795 | -106.1493
17.3831 | -106.148
17.3867 | -106.1466
17.3904 | -106.1444
17.394 | -106.1408
17.3976 | -106.1391
17.4012 | -106.1372
17.4049 | -106.1359
17.4085 | -106.1335
17.4121 | -106.1291
17.4158 | -106.1266
17.4194 | -106.1249
17.423 | -106.1209
17.4266 | -106.117
17.4303 | -106.1141
17.4339 | -106.1095
17.4375 | -106.106
17.4412 | -106.0999
17.4448 | -106.0957
17.4484 | -106.089
17.452 | -106.0842
17.4557 | -106.0775
17.4593 | -106.0703
17.4629 | -106.0657
17.4666 | -106.0613
17.4702 | -106.0576
17.4738 | -106.0518
17.4774 | -106.0442
17.4811 | -106.0363
17.4847 | -106.0259
17.4883 | -106.0165
17.4919 | -106.0067
17.4956 | -105.9981
17.4992 | -105.9886
17.5028 | -105.9794
17.5065 | -105.9689
17.5101 | -105.9584
17.5137 | -105.9497
17.5173 | -105.9395
17.521 | -105.9299
17.5246 | -105.919
17.5282 | -105.9066
17.5319 | -105.8955
17.5355 | -105.8864
17.5391 | -105.8776
17.5427 | -105.8681
17.5464 | -105.8594
17.55 | -105.8533
17.5536 | -105.8493
17.5573 | -105.8445
17.5609 | -105.8384
17.5645 | -105.8304
17.5681 | -105.8244
17.5718 | -105.8189
17.5754 | -105.8142
17.579 | -105.8095
17.5827 | -105.8054
17.5863 | -105.8051
17.5899 | -105.8049
17.5935 | -105.8034
17.5972 | -105.8032
17.6008 | -105.8032
17.6044 | -105.8037
17.6081 | -105.806
17.6117 | -105.8078
17.6153 | -105.8071
17.6189 | -105.8077
17.6226 | -105.8085
17.6262 | -105.81
17.6298 | -105.8125
17.6335 | -105.8137
17.6371 | -105.813
17.6407 | -105.816
17.6443 | -105.8199
17.648 | -105.8237
17.6516 | -105.8265
17.6552 | -105.8277
17.6589 | -105.8306
17.6625 | -105.8363
17.6661 | -105.8421
17.6697 | -105.8448
17.6734 | -105.8472
17.677 | -105.8493
17.6806 | -105.8531
17.6842 | -105.8586
17.6879 | -105.8651
17.6915 | -105.8733
17.6951 | -105.88
17.6988 | -105.8888
17.7024 | -105.8942
17.706 | -105.899
17.7096 | -105.9067
17.7133 | -105.9134
17.7169 | -105.9195
17.7205 | -105.9269
17.7242 | -105.9337
17.7278 | -105.9394
17.7314 | -105.9456
17.735 | -105.9548
17.7387 | -105.961
17.7423 | -105.967
17.7459 | -105.9716
17.7496 | -105.9762
17.7532 | -105.9817
17.7568 | -105.9883
17.7604 | -105.9932
17.7641 | -105.9972
17.7677 | -106.0014
17.7713 | -106.0049
17.775 | -106.0105
17.7786 | -106.0165
17.7822 | -106.0153
17.7858 | -106.0129
17.7895 | -106.0122
17.7931 | -106.0134
17.7967 | -106.0102
17.8004 | -106.0112
17.804 | -106.0121
17.8076 | -106.0116
17.8112 | -106.0123
17.8149 | -106.0152
17.8185 | -106.0189
17.8221 | -106.0181
17.8258 | -106.0159
17.8294 | -106.0154
17.833 | -106.0145
17.8366 | -106.015
17.8403 | -106.0143
17.8439 | -106.0123
17.8475 | -106.0112
17.8512 | -106.0088
17.8548 | -106.0076
17.8584 | -106.0081
17.862 | -106.0071
17.8657 | -106.004
17.8693 | -106.0024
17.8729 | -106.0018
17.8765 | -106.0029
17.8802 | -106.0033
17.8838 | -106.0003
17.8874 | -106.0005
17.8911 | -105.999
17.8947 | -105.9935
17.8983 | -105.9881
17.9019 | -105.9829
17.9056 | -105.9766
17.9092 | -105.971
17.9128 | -105.9648
17.9165 | -105.9603
17.9201 | -105.9547
17.9237 | -105.9501
17.9273 | -105.9489
17.931 | -105.9465
17.9346 | -105.9403
17.9382 | -105.9328
17.9419 | -105.926
17.9455 | -105.9184
17.9491 | -105.9127
17.9527 | -105.9089
17.9564 | -105.9061
17.96 | -105.905
17.9636 | -105.903
17.9673 | -105.9004
17.9709 | -105.8973
17.9745 | -105.895
17.9781 | -105.896
17.9818 | -105.8972
17.9854 | -105.8978
17.989 | -105.8979
17.9927 | -105.9026
17.9963 | -105.9073
17.9999 | -105.9108
18.0035 | -105.9111
18.0072 | -105.9111
18.0108 | -105.9146
18.0144 | -105.9191
18.0181 | -105.923
18.0217 | -105.9267
18.0253 | -105.9323
18.0289 | -105.9377
18.0326 | -105.943
18.0362 | -105.9491
18.0398 | -105.9547
18.0434 | -105.9588
18.0471 | -105.9617
18.0507 | -105.9636
18.0543 | -105.9677
18.058 | -105.9677
18.0616 | -105.9635
18.0652 | -105.9606
18.0688 | -105.9584
18.0725 | -105.956
18.0761 | -105.9552
18.0797 | -105.9567
18.0834 | -105.9571
18.087 | -105.9535
18.0906 | -105.9495
18.0942 | -105.9442
18.0979 | -105.9385
18.1015 | -105.9318
18.1051 | -105.9259
18.1088 | -105.9212
18.1124 | -105.9143
18.116 | -105.9063
18.1196 | -105.8993
18.1233 | -105.8939
18.1269 | -105.8893
18.1305 | -105.8837
18.1342 | -105.8768
18.1378 | -105.8705
18.1414 | -105.866
18.145 | -105.8612
18.1487 | -105.8572
18.1523 | -105.855
18.1559 | -105.8528
18.1596 | -105.8503
18.1632 | -105.8515
18.1668 | -105.8544
18.1704 | -105.8564
18.1741 | -105.855
18.1777 | -105.8568
18.1813 | -105.8584
18.185 | -105.8596
18.1886 | -105.8619
18.1922 | -105.8641
18.1958 | -105.8683
18.1995 | -105.872
18.2031 | -105.8752
18.2067 | -105.8796
18.2103 | -105.8829
18.214 | -105.8885
18.2176 | -105.8932
18.2212 | -105.8981
18.2249 | -105.9026
18.2285 | -105.9094
18.2321 | -105.9169
18.2357 | -105.9217
18.2394 | -105.9276
18.243 | -105.9336
18.2466 | -105.9394
18.2503 | -105.9429
18.2539 | -105.947
18.2575 | -105.952
18.2611 | -105.957
18.2648 | -105.9615
18.2684 | -105.9632
18.272 | -105.9647
18.2757 | -105.9653
18.2793 | -105.9668
18.2829 | -105.9633
18.2865 | -105.958
18.2902 | -105.956
18.2938 | -105.9545
18.2974 | -105.9494
18.3011 | -105.9449
18.3047 | -105.942
18.3083 | -105.9375
18.3119 | -105.9302
18.3156 | -105.9235
18.3192 | -105.9149
18.3228 | -105.9103
18.3265 | -105.9039
18.3301 | -105.8967
18.3337 | -105.8874
18.3373 | -105.8755
18.341 | -105.8643
18.3446 | -105.8561
18.3482 | -105.8484
18.3519 | -105.8367
18.3555 | -105.8233
18.3591 | -105.8123
18.3627 | -105.8015
18.3664 | -105.789
18.37 | -105.7776
18.3736 | -105.7669
18.3773 | -105.7545
18.3809 | -105.7425
18.3845 | -105.7315
18.3881 | -105.7169
18.3918 | -105.7036
18.3954 | -105.6909
18.399 | -105.6773
18.4026 | -105.6645
18.4063 | -105.6514
18.4099 | -105.6387
18.4135 | -105.6273
18.4172 | -105.6172
18.4208 | -105.6058
18.4244 | -105.5927
18.428 | -105.5784
18.4317 | -105.5678
18.4353 | -105.5603
18.4389 | -105.5505
18.4426 | -105.5435
18.4462 | -105.5364
18.4498 | -105.5302
18.4534 | -105.5254
18.4571 | -105.5214
18.4607 | -105.5178
18.4643 | -105.5163
18.468 | -105.514
18.4716 | -105.5144
18.4752 | -105.5154
18.4788 | -105.5153
18.4825 | -105.5148
18.4861 | -105.5163
18.4897 | -105.5186
18.4934 | -105.521
18.497 | -105.5253
18.5006 | -105.5287
18.5042 | -105.5322
18.5079 | -105.5361
18.5115 | -105.5409
18.5151 | -105.5463
18.5188 | -105.5546
18.5224 | -105.5626
18.526 | -105.5701
18.5296 | -105.5785
18.5333 | -105.5856
18.5369 | -105.595
18.5405 | -105.6021
18.5442 | -105.6103
18.5478 | -105.6171
18.5514 | -105.6228
18.555 | -105.6288
18.5587 | -105.6364
18.5623 | -105.6429
18.5659 | -105.6469
18.5695 | -105.6532
18.5732 | -105.6617
18.5768 | -105.6681
18.5804 | -105.6737
18.5841 | -105.6793
18.5877 | -105.6856
18.5913 | -105.6905
18.5949 | -105.6955
18.5986 | -105.7006
18.6022 | -105.7058
18.6058 | -105.7101
18.6095 | -105.7121
18.6131 | -105.7125
18.6167 | -105.7153
18.6203 | -105.719
18.624 | -105.7205
18.6276 | -105.7201
18.6312 | -105.7192
18.6349 | -105.7159
18.6385 | -105.7116
18.6421 | -105.7091
18.6457 | -105.7051
18.6494 | -105.6988
18.653 | -105.695
18.6566 | -105.6906
18.6603 | -105.6833
18.6639 | -105.6773
18.6675 | -105.672
18.6711 | -105.6662
18.6748 | -105.6609
18.6784 | -105.6548
18.682 | -105.6475
18.6857 | -105.6426
18.6893 | -105.6361
18.6929 | -105.6309
18.6965 | -105.625
18.7002 | -105.6201
18.7038 | -105.6146
18.7074 | -105.6089
18.7111 | -105.6044
18.7147 | -105.5995
18.7183 | -105.597
18.7219 | -105.5933
18.7256 | -105.5894
18.7292 | -105.5856
18.7328 | -105.5808
18.7365 | -105.5769
18.7401 | -105.5745
18.7437 | -105.574
18.7473 | -105.5731
18.751 | -105.5741
18.7546 | -105.5751
18.7582 | -105.5754
18.7618 | -105.576
18.7655 | -105.5784
18.7691 | -105.5805
18.7727 | -105.5835
18.7764 | -105.5871
18.78 | -105.5894
18.7836 | -105.5899
18.7872 | -105.5922
18.7909 | -105.5947
18.7945 | -105.5964
18.7981 | -105.6007
18.8018 | -105.6057
18.8054 | -105.611
18.809 | -105.6164
18.8126 | -105.6231
18.8163 | -105.6279
18.8199 | -105.6331
18.8235 | -105.6361
18.8272 | -105.6368
18.8308 | -105.6414
18.8344 | -105.645
18.838 | -105.6445
18.8417 | -105.6458
18.8453 | -105.6494
18.8489 | -105.6518
18.8526 | -105.6538
18.8562 | -105.6563
18.8598 | -105.6591
18.8634 | -105.6604
18.8671 | -105.6625
18.8707 | -105.6644
18.8743 | -105.6665
18.878 | -105.6693
18.8816 | -105.6726
18.8852 | -105.6776
18.8888 | -105.6791
18.8925 | -105.6813
18.8961 | -105.687
18.8997 | -105.6916
18.9034 | -105.6962
18.907 | -105.7005
18.9106 | -105.704
18.9142 | -105.7074
18.9179 | -105.7124
18.9215 | -105.7163
18.9251 | -105.7203
18.9288 | -105.7277
18.9324 | -105.7328
18.936 | -105.7366
18.9396 | -105.7415
18.9433 | -105.7461
18.9469 | -105.7519
18.9505 | -105.7577
18.9541 | -105.7631
18.9578 | -105.7718
18.9614 | -105.7783
18.965 | -105.7823
18.9687 | -105.7867
18.9723 | -105.794
18.9759 | -105.8
18.9795 | -105.8026
18.9832 | -105.8048
18.9868 | -105.8074
18.9904 | -105.8101
18.9941 | -105.8133
18.9977 | -105.8152
19.0013 | -105.8182
19.0049 | -105.8177
19.0086 | -105.8158
19.0122 | -105.8121
19.0158 | -105.8086
19.0195 | -105.8049
19.0231 | -105.7984
19.0267 | -105.7907
19.0303 | -105.7848
19.034 | -105.7796
19.0376 | -105.7742
19.0412 | -105.7665
19.0449 | -105.7573
19.0485 | -105.7492
19.0521 | -105.7434
19.0557 | -105.7379
19.0594 | -105.7331
19.063 | -105.7291
19.0666 | -105.7232
19.0703 | -105.7141
19.0739 | -105.7049
19.0775 | -105.6983
19.0811 | -105.6931
19.0848 | -105.6881
19.0884 | -105.6834
19.092 | -105.6805
19.0957 | -105.6782
19.0993 | -105.674
19.1029 | -105.6693
19.1065 | -105.6687
19.1102 | -105.6686
19.1138 | -105.6674
19.1174 | -105.6673
19.121 | -105.6675
19.1247 | -105.667
19.1283 | -105.6665
19.1319 | -105.6674
19.1356 | -105.6697
19.1392 | -105.6713
19.1428 | -105.6721
19.1464 | -105.671
19.1501 | -105.6704
19.1537 | -105.6718
19.1573 | -105.6708
19.161 | -105.6719
19.1646 | -105.6727
19.1682 | -105.6707
19.1718 | -105.6683
19.1755 | -105.6652
19.1791 | -105.6623
19.1827 | -105.6601
19.1864 | -105.6577
19.19 | -105.6539
19.1936 | -105.6496
19.1972 | -105.6432
19.2009 | -105.6344
19.2045 | -105.6265
19.2081 | -105.619
19.2118 | -105.6118
19.2154 | -105.6037
19.219 | -105.5919
19.2226 | -105.5807
19.2263 | -105.57
19.2299 | -105.558
19.2335 | -105.5456
19.2372 | -105.5327
19.2408 | -105.5185
19.2444 | -105.5044
19.248 | -105.4913
19.2517 | -105.4808
19.2553 | -105.4705
19.2589 | -105.4595
19.2626 | -105.449
19.2662 | -105.4375
19.2698 | -105.4271
19.2734 | -105.4168
19.2771 | -105.4051
19.2807 | -105.3938
19.2843 | -105.3845
19.2879 | -105.3744
19.2916 | -105.3631
19.2952 | -105.3546
19.2988 | -105.3479
19.3025 | -105.3381
19.3061 | -105.3283
19.3097 | -105.3199
19.3133 | -105.3124
19.317 | -105.3059
19.3206 | -105.3001
19.3242 | -105.2937
19.3279 | -105.2859
19.3315 | -105.2785
19.3351 | -105.2735
19.3387 | -105.2704
19.3424 | -105.2693
19.346 | -105.2688
19.3496 | -105.267
19.3533 | -105.2633
19.3569 | -105.2609
19.3605 | -105.2585
19.3641 | -105.257
19.3678 | -105.254
19.3714 | -105.2523
19.375 | -105.2514
19.3787 | -105.2491
19.3823 | -105.2455
19.3859 | -105.241
19.3895 | -105.2378
19.3932 | -105.2346
19.3968 | -105.2313
19.4004 | -105.2285
19.4041 | -105.2248
19.4077 | -105.222
19.4113 | -105.2184
19.4149 | -105.2163
19.4186 | -105.213
19.4222 | -105.2091
19.4258 | -105.2062
19.4295 | -105.2034
19.4331 | -105.2015
19.4367 | -105.1967
19.4403 | -105.1887
19.444 | -105.1838
19.4476 | -105.1809
19.4512 | -105.1792
19.4549 | -105.178
19.4585 | -105.1746
19.4621 | -105.1715
19.4657 | -105.1684
19.4694 | -105.1658
19.473 | -105.1639
19.4766 | -105.1644
19.4802 | -105.1667
19.4839 | -105.1651
19.4875 | -105.1631
19.4911 | -105.1641
19.4948 | -105.1666
19.4984 | -105.1687
19.502 | -105.1708
19.5056 | -105.1717
19.5093 | -105.1735
19.5129 | -105.1758
19.5165 | -105.1787
19.5202 | -105.1806
19.5238 | -105.1796
19.5274 | -105.1775
19.531 | -105.1782
19.5347 | -105.1766
19.5383 | -105.1719
19.5419 | -105.1691
19.5456 | -105.1688
19.5492 | -105.1699
19.5528 | -105.1722
19.5564 | -105.1731
19.5601 | -105.1722
19.5637 | -105.1722
19.5673 | -105.172
19.571 | -105.1702
19.5746 | -105.1686
19.5782 | -105.1677
19.5818 | -105.1646
19.5855 | -105.1608
19.5891 | -105.1582
19.5927 | -105.1546
19.5964 | -105.1509
19.6 | -105.1481
19.6036 | -105.1463
19.6072 | -105.1436
19.6109 | -105.1394
19.6145 | -105.1352
19.6181 | -105.1314
19.6218 | -105.1294
19.6254 | -105.1267
19.629 | -105.1252
19.6326 | -105.124
19.6363 | -105.1257
19.6399 | -105.1277
19.6435 | -105.1293
19.6471 | -105.1345
19.6508 | -105.1388
19.6544 | -105.1402
19.658 | -105.145
19.6617 | -105.1507
19.6653 | -105.1553
19.6689 | -105.1591
19.6725 | -105.1645
19.6762 | -105.1717
19.6798 | -105.1808
19.6834 | -105.1897
19.6871 | -105.1982
19.6907 | -105.2082
19.6943 | -105.2151
19.6979 | -105.2245
19.7016 | -105.2373
19.7052 | -105.2484
19.7088 | -105.2579
19.7125 | -105.2681
19.7161 | -105.2784
19.7197 | -105.2884
19.7233 | -105.2981
19.727 | -105.3078
19.7306 | -105.3146
19.7342 | -105.3205
19.7379 | -105.3264
19.7415 | -105.3334
19.7451 | -105.3402
19.7487 | -105.3435
19.7524 | -105.3446
19.756 | -105.3442
19.7596 | -105.3422
19.7633 | -105.3413
19.7669 | -105.3383
19.7705 | -105.3323
19.7741 | -105.3268
19.7778 | -105.3205
19.7814 | -105.3134
19.785 | -105.3074
19.7887 | -105.3024
19.7923 | -105.2971
19.7959 | -105.2907
19.7995 | -105.283
19.8032 | -105.2761
19.8068 | -105.2689
19.8104 | -105.2587
19.8141 | -105.252
19.8177 | -105.2474
19.8213 | -105.2405
19.8249 | -105.2309
19.8286 | -105.2211
19.8322 | -105.2141
19.8358 | -105.209
19.8394 | -105.2055
19.8431 | -105.2018
19.8467 | -105.1968
19.8503 | -105.193
19.854 | -105.1905
19.8576 | -105.1912
19.8612 | -105.1931
19.8648 | -105.1939
19.8685 | -105.1963
19.8721 | -105.1998
19.8757 | -105.2047
19.8794 | -105.2081
19.883 | -105.2114
19.8866 | -105.2183
19.8902 | -105.2292
19.8939 | -105.2359
19.8975 | -105.2445
19.9011 | -105.2534
19.9048 | -105.2615
19.9084 | -105.2715
19.912 | -105.2817
19.9156 | -105.2917
19.9193 | -105.2993
19.9229 | -105.3067
19.9265 | -105.3165
19.9302 | -105.3241
19.9338 | -105.3339
19.9374 | -105.3465
19.941 | -105.3572
19.9447 | -105.3656
19.9483 | -105.3766
19.9519 | -105.3849
19.9556 | -105.3919
19.9592 | -105.3981
19.9628 | -105.4026
19.9664 | -105.4093
19.9701 | -105.4145
19.9737 | -105.4193
19.9773 | -105.4256
19.981 | -105.4297
19.9846 | -105.4327
19.9882 | -105.4354
19.9918 | -105.4397
19.9955 | -105.443
19.9991 | -105.4475
20.0027 | -105.4543
20.0063 | -105.4589
20.01 | -105.4606
20.0136 | -105.4635
20.0172 | -105.4655
20.0209 | -105.4675
20.0245 | -105.4698
20.0281 | -105.4719
20.0317 | -105.474
20.0354 | -105.4762
20.039 | -105.4793
20.0426 | -105.4812
20.0463 | -105.4848
20.0499 | -105.4895
20.0535 | -105.4942
20.0571 | -105.4976
20.0608 | -105.4998
20.0644 | -105.504
20.068 | -105.506
20.0717 | -105.5066
20.0753 | -105.5055
20.0789 | -105.5044
20.0825 | -105.504
20.0862 | -105.5018
20.0898 | -105.4992
20.0934 | -105.4978
20.0971 | -105.4943
20.1007 | -105.4905
20.1043 | -105.4863
20.1079 | -105.4825
20.1116 | -105.4779
20.1152 | -105.4709
20.1188 | -105.4634
20.1225 | -105.4548
20.1261 | -105.4465
20.1297 | -105.4394
20.1333 | -105.4336
20.137 | -105.4268
20.1406 | -105.4172
20.1442 | -105.4068
20.1479 | -105.3984
20.1515 | -105.3891
20.1551 | -105.3797
20.1587 | -105.3697
20.1624 | -105.3609
20.166 | -105.3539
20.1696 | -105.3426
20.1732 | -105.3297
20.1769 | -105.3198
20.1805 | -105.309
20.1841 | -105.2972
20.1878 | -105.2888
20.1914 | -105.2813
20.195 | -105.2722
20.1986 | -105.2646
20.2023 | -105.2563
20.2059 | -105.2473
20.2095 | -105.2406
20.2132 | -105.2325
20.2168 | -105.2235
20.2204 | -105.2152
20.224 | -105.2077
20.2277 | -105.2014
20.2313 | -105.1964
20.2349 | -105.19
20.2386 | -105.1842
20.2422 | -105.1776
20.2458 | -105.1714
20.2494 | -105.1653
20.2531 | -105.1583
20.2567 | -105.1547
20.2603 | -105.1512
20.264 | -105.1478
20.2676 | -105.1445
20.2712 | -105.1387
20.2748 | -105.1338
20.2785 | -105.1304
20.2821 | -105.1275
20.2857 | -105.1263
20.2894 | -105.125
20.293 | -105.1263
20.2966 | -105.1279
20.3002 | -105.1285
20.3039 | -105.1297
20.3075 | -105.1327
20.3111 | -105.1345
20.3148 | -105.1375
20.3184 | -105.1422
20.322 | -105.1465
20.3256 | -105.1486
20.3293 | -105.1496
20.3329 | -105.1508
20.3365 | -105.1517
20.3402 | -105.1515
20.3438 | -105.1523
20.3474 | -105.1541
20.351 | -105.153
20.3547 | -105.1529
20.3583 | -105.1533
20.3619 | -105.1537
20.3655 | -105.1533
20.3692 | -105.1537
20.3728 | -105.1562
20.3764 | -105.1572
20.3801 | -105.1562
20.3837 | -105.1549
20.3873 | -105.1527
20.3909 | -105.151
20.3946 | -105.1478
20.3982 | -105.1453
20.4018 | -105.1423
20.4055 | -105.1378
20.4091 | -105.1348
20.4127 | -105.1307
20.4163 | -105.1294
20.42 | -105.1276
20.4236 | -105.1243
20.4272 | -105.1196
20.4309 | -105.1163
20.4345 | -105.1167
20.4381 | -105.1162
20.4417 | -105.1162
20.4454 | -105.1167
20.449 | -105.1172
20.4526 | -105.1172
20.4563 | -105.1157
20.4599 | -105.1171
20.4635 | -105.1195
20.4671 | -105.1219
20.4708 | -105.1245
20.4744 | -105.1253
20.478 | -105.1214
20.4817 | -105.12
20.4853 | -105.1212
20.4889 | -105.1209
20.4925 | -105.1223
20.4962 | -105.1255
20.4998 | -105.1269
20.5034 | -105.1285
20.5071 | -105.1322
20.5107 | -105.1333
20.5143 | -105.1344
20.5179 | -105.1352
20.5216 | -105.1358
20.5252 | -105.1358
20.5288 | -105.1366
20.5325 | -105.1383
20.5361 | -105.1404
20.5397 | -105.1422
20.5433 | -105.1417
20.547 | -105.1391
20.5506 | -105.137
20.5542 | -105.1373
20.5578 | -105.1386
20.5615 | -105.1405
20.5651 | -105.1403
20.5687 | -105.139
20.5724 | -105.1383
20.576 | -105.1364
20.5796 | -105.133
20.5832 | -105.1276
20.5869 | -105.1236
20.5905 | -105.1214
20.5941 | -105.1203
20.5978 | -105.119
20.6014 | -105.1168
20.605 | -105.112
20.6086 | -105.1061
20.6123 | -105.1023
20.6159 | -105.098
20.6195 | -105.0942
20.6232 | -105.0907
20.6268 | -105.0861
20.6304 | -105.0815
20.634 | -105.0761
20.6377 | -105.0709
20.6413 | -105.0653
20.6449 | -105.0601
20.6486 | -105.0532
20.6522 | -105.048
20.6558 | -105.0427
20.6594 | -105.0362
20.6631 | -105.0297
20.6667 | -105.023
20.6703 | -105.016
20.674 | -105.0092
20.6776 | -105.0037
20.6812 | -104.9982
20.6848 | -104.9921
20.6885 | -104.9854
20.6921 | -104.9815
20.6957 | -104.9779
20.6994 | -104.9726
20.703 | -104.9668
20.7066 | -104.9636
20.7102 | -104.9593
20.7139 | -104.9546
20.7175 | -104.9509
20.7211 | -104.9477
20.7247 | -104.9449
20.7284 | -104.9403
20.732 | -104.9374
20.7356 | -104.9337
20.7393 | -104.9313
20.7429 | -104.9311
20.7465 | -104.9319
20.7501 | -104.9344
20.7538 | -104.9351
20.7574 | -104.9349
20.761 | -104.9355
20.7647 | -104.9357
20.7683 | -104.9381
20.7719 | -104.9419
20.7755 | -104.9435
20.7792 | -104.945
20.7828 | -104.9467
20.7864 | -104.9507
20.7901 | -104.954
20.7937 | -104.9595
20.7973 | -104.9655
20.8009 | -104.9708
20.8046 | -104.9756
20.8082 | -104.9814
20.8118 | -104.9887
20.8155 | -104.995
20.8191 | -105.0025
20.8227 | -105.0107
20.8263 | -105.0197
20.83 | -105.0276
20.8336 | -105.037
20.8372 | -105.0469
20.8409 | -105.0568
20.8445 | -105.0664
20.8481 | -105.074
20.8517 | -105.0793
20.8554 | -105.0855
20.859 | -105.096
20.8626 | -105.106
20.8663 | -105.1156
20.8699 | -105.124
20.8735 | -105.1334
20.8771 | -105.1444
20.8808 | -105.1547
20.8844 | -105.1636
20.888 | -105.1679
20.8916 | -105.1715
20.8953 | -105.1779
20.8989 | -105.1863
20.9025 | -105.196
20.9062 | -105.2027
20.9098 | -105.2081
20.9134 | -105.2144
20.917 | -105.2188
20.9207 | -105.2249
20.9243 | -105.2304
20.9279 | -105.2328
20.9316 | -105.235
20.9352 | -105.2321
20.9388 | -105.2337
20.9424 | -105.2353
20.9461 | -105.2353
20.9497 | -105.2334
20.9533 | -105.231
20.957 | -105.229
20.9606 | -105.2254
20.9642 | -105.2218
20.9678 | -105.2184
20.9715 | -105.2117
20.9751 | -105.2052
20.9787 | -105.1993
20.9824 | -105.194
20.986 | -105.1877
20.9896 | -105.1808
20.9932 | -105.1737
20.9969 | -105.1671
21.0005 | -105.1594
21.0041 | -105.1486
21.0078 | -105.1391
21.0114 | -105.1305
21.015 | -105.1186
21.0186 | -105.106
21.0223 | -105.093
21.0259 | -105.0824
21.0295 | -105.0702
21.0332 | -105.0582
21.0368 | -105.0473
21.0404 | -105.035
21.044 | -105.0233
21.0477 | -105.011
21.0513 | -105.0012
21.0549 | -104.9926
21.0586 | -104.984
21.0622 | -104.9747
21.0658 | -104.9637
21.0694 | -104.9525
21.0731 | -104.9432
21.0767 | -104.934
21.0803 | -104.9265
21.0839 | -104.9156
21.0876 | -104.906
21.0912 | -104.8949
21.0948 | -104.8841
21.0985 | -104.8757
21.1021 | -104.8709
21.1057 | -104.865
21.1093 | -104.8591
21.113 | -104.8541
21.1166 | -104.8506
21.1202 | -104.8473
21.1239 | -104.8433
21.1275 | -104.8397
21.1311 | -104.8364
21.1347 | -104.8343
21.1384 | -104.8329
21.142 | -104.8336
21.1456 | -104.8352
21.1493 | -104.8394
21.1529 | -104.8422
21.1565 | -104.8419
21.1601 | -104.8433
21.1638 | -104.8462
21.1674 | -104.8514
21.171 | -104.8558
21.1747 | -104.8583
21.1783 | -104.8611
21.1819 | -104.8649
21.1855 | -104.8732
21.1892 | -104.8807
21.1928 | -104.8854
21.1964 | -104.8905
21.2001 | -104.8966
21.2037 | -104.9024
21.2073 | -104.9087
21.2109 | -104.9151
21.2146 | -104.9235
21.2182 | -104.9312
21.2218 | -104.9381
21.2255 | -104.9478
21.2291 | -104.9567
21.2327 | -104.965
21.2363 | -104.9726
21.24 | -104.98
21.2436 | -104.9856
21.2472 | -104.993
21.2508 | -105.0013
21.2545 | -105.0081
21.2581 | -105.0148
21.2617 | -105.021
21.2654 | -105.0273
21.269 | -105.0312
21.2726 | -105.0352
21.2762 | -105.038
21.2799 | -105.0388
21.2835 | -105.0398
21.2871 | -105.0402
21.2908 | -105.0397
21.2944 | -105.0408
21.298 | -105.038
21.3016 | -105.0353
21.3053 | -105.0343
21.3089 | -105.0325
21.3125 | -105.0305
21.3162 | -105.0294
21.3198 | -105.029
21.3234 | -105.0296
21.327 | -105.0285
21.3307 | -105.0271
21.3343 | -105.0263
21.3379 | -105.0242
21.3416 | -105.0232
21.3452 | -105.0225
21.3488 | -105.0216
21.3524 | -105.0222
21.3561 | -105.0228
21.3597 | -105.0243
21.3633 | -105.0266
21.367 | -105.0302
21.3706 | -105.0349
21.3742 | -105.0387
21.3778 | -105.0397
21.3815 | -105.0431
21.3851 | -105.0499
21.3887 | -105.0567
21.3924 | -105.0659
21.396 | -105.0754
21.3996 | -105.0846
21.4032 | -105.0921
21.4069 | -105.1006
21.4105 | -105.109
21.4141 | -105.1193
21.4178 | -105.1293
21.4214 | -105.138
21.425 | -105.1478
21.4286 | -105.1586
21.4323 | -105.1697
21.4359 | -105.1813
21.4395 | -105.1899
21.4431 | -105.1973
21.4468 | -105.2059
21.4504 | -105.2158
21.454 | -105.2263
21.4577 | -105.2367
21.4613 | -105.2431
21.4649 | -105.2519
21.4685 | -105.2585
21.4722 | -105.2628
21.4758 | -105.2682
21.4794 | -105.2744
21.4831 | -105.2801
21.4867 | -105.2832
21.4903 | -105.2839
21.4939 | -105.2861
21.4976 | -105.2887
21.5012 | -105.2889
21.5048 | -105.2889
21.5085 | -105.2885
21.5121 | -105.2881
21.5157 | -105.2851
21.5193 | -105.279
21.523 | -105.2728
21.5266 | -105.2687
21.5302 | -105.2617
21.5339 | -105.2513
21.5375 | -105.2415
21.5411 | -105.2327
21.5447 | -105.2237
21.5484 | -105.2138
21.552 | -105.2026
21.5556 | -105.1878
21.5593 | -105.1758
21.5629 | -105.1626
21.5665 | -105.1491
21.5701 | -105.1361
21.5738 | -105.1244
21.5774 | -105.1145
21.581 | -105.1037
21.5847 | -105.0928
21.5883 | -105.0815
21.5919 | -105.0691
21.5955 | -105.0561
21.5992 | -105.044
21.6028 | -105.0368
21.6064 | -105.0299
21.6101 | -105.0209
21.6137 | -105.0113
21.6173 | -105.0039
21.6209 | -104.9978
21.6246 | -104.9922
21.6282 | -104.9866
21.6318 | -104.9781
21.6354 | -104.9723
21.6391 | -104.9688
21.6427 | -104.9652
21.6463 | -104.9639
21.65 | -104.9642
21.6536 | -104.9658
21.6572 | -104.9639
21.6608 | -104.9633
21.6645 | -104.9631
21.6681 | -104.9664
21.6717 | -104.9703
21.6754 | -104.9729
21.679 | -104.977
21.6826 | -104.9835
21.6862 | -104.9868
21.6899 | -104.9878
21.6935 | -104.9896
21.6971 | -104.9903
21.7008 | -104.9921
21.7044 | -104.9945
21.708 | -104.9979
21.7116 | -105.0012
21.7153 | -105.0038
21.7189 | -105.0046
21.7225 | -105.0023
21.7262 | -105.0034
21.7298 | -105.0054
21.7334 | -105.0059
21.737 | -105.0043
21.7407 | -105.0029
21.7443 | -105.0042
21.7479 | -105.0047
21.7516 | -105.0036
21.7552 | -105.0016
21.7588 | -104.9999
21.7624 | -104.9982
21.7661 | -104.996
21.7697 | -104.9932
21.7733 | -104.99
21.777 | -104.9884
21.7806 | -104.9884
21.7842 | -104.9852
21.7878 | -104.9807
21.7915 | -104.9762
21.7951 | -104.9706
21.7987 | -104.9646
21.8023 | -104.956
21.806 | -104.9482
21.8096 | -104.9406
21.8132 | -104.9303
21.8169 | -104.9207
21.8205 | -104.9112
21.8241 | -104.9022
21.8277 | -104.8953
21.8314 | -104.8874
21.835 | -104.878
21.8386 | -104.8692
21.8423 | -104.8617
21.8459 | -104.8531
21.8495 | -104.8414
21.8531 | -104.8286
21.8568 | -104.8188
21.8604 | -104.8076
21.864 | -104.7989
21.8677 | -104.7894
21.8713 | -104.7795
21.8749 | -104.7684
21.8785 | -104.756
21.8822 | -104.7438
21.8858 | -104.7294
21.8894 | -104.7149
21.8931 | -104.7006
21.8967 | -104.6874
21.9003 | -104.6761
21.9039 | -104.6644
21.9076 | -104.6515
21.9112 | -104.6362
21.9148 | -104.6214
21.9185 | -104.6078
21.9221 | -104.5961
21.9257 | -104.5814
21.9293 | -104.5667
21.933 | -104.5533
21.9366 | -104.5381
21.9402 | -104.5251
21.9439 | -104.51
21.9475 | -104.4948
21.9511 | -104.4816
21.9547 | -104.4653
21.9584 | -104.4471
21.962 | -104.4309
21.9656 | -104.4183
21.9692 | -104.4054
21.9729 | -104.3922
21.9765 | -104.3777
21.9801 | -104.3637
21.9838 | -104.3501
21.9874 | -104.3345
21.991 | -104.3195
21.9946 | -104.3056
21.9983 | -104.2923
22.0019 | -104.281
22.0055 | -104.2687
22.0092 | -104.2572
22.0128 | -104.2451
22.0164 | -104.2337
22.02 | -104.2227
22.0237 | -104.2112
22.0273 | -104.2014
22.0309 | -104.1924
22.0346 | -104.1867
22.0382 | -104.1798
22.0418 | -104.1718
22.0454 | -104.1647
22.0491 | -104.1597
22.0527 | -104.155
22.0563 | -104.1508
22.06 | -104.1471
22.0636 | -104.1448
22.0672 | -104.1426
22.0708 | -104.1401
22.0745 | -104.1396
22.0781 | -104.1393
22.0817 | -104.1385
22.0854 | -104.1393
22.089 | -104.1411
22.0926 | -104.1436
22.0962 | -104.1491
22.0999 | -104.1562
22.1035 | -104.165
22.1071 | -104.1761
22.1108 | -104.1863
22.1144 | -104.1963
22.118 | -104.2049
22.1216 | -104.215
22.1253 | -104.2242
22.1289 | -104.2352
22.1325 | -104.2469
22.1362 | -104.2587
22.1398 | -104.27
22.1434 | -104.2829
22.147 | -104.2962
22.1507 | -104.3097
22.1543 | -104.3244
22.1579 | -104.3387
22.1615 | -104.3552
22.1652 | -104.3715
22.1688 | -104.3842
22.1724 | -104.3975
22.1761 | -104.4112
22.1797 | -104.424
22.1833 | -104.4359
22.1869 | -104.45
22.1906 | -104.4659
22.1942 | -104.4813
22.1978 | -104.4955
22.2015 | -104.5085
22.2051 | -104.5221
22.2087 | -104.5327
22.2123 | -104.5432
22.216 | -104.5538
22.2196 | -104.5636
22.2232 | -104.5742
22.2269 | -104.5842
22.2305 | -104.5933
22.2341 | -104.5995
22.2377 | -104.6073
22.2414 | -104.6157
22.245 | -104.6228
22.2486 | -104.6294
22.2523 | -104.6336
22.2559 | -104.6385
22.2595 | -104.6422
22.2631 | -104.6461
22.2668 | -104.6524
22.2704 | -104.6588
22.274 | -104.6622
22.2777 | -104.6615
22.2813 | -104.6613
22.2849 | -104.6617
22.2885 | -104.6625
22.2922 | -104.6633
22.2958 | -104.6646
22.2994 | -104.6656
22.3031 | -104.6685
22.3067 | -104.6697
22.3103 | -104.6696
22.3139 | -104.6701
22.3176 | -104.6691
22.3212 | -104.6691
22.3248 | -104.6676
22.3285 | -104.6688
22.3321 | -104.6692
22.3357 | -104.6677
22.3393 | -104.6676
22.343 | -104.6671
22.3466 | -104.667
22.3502 | -104.6665
22.3538 | -104.6657
22.3575 | -104.6637
22.3611 | -104.6624
22.3647 | -104.6618
22.3684 | -104.6612
22.372 | -104.6601
22.3756 | -104.6593
22.3792 | -104.6592
22.3829 | -104.6598
22.3865 | -104.6616
22.3901 | -104.6605
22.3938 | -104.661
22.3974 | -104.6644
22.401 | -104.6652
22.4046 | -104.6658
22.4083 | -104.6669
22.4119 | -104.6672
22.4155 | -104.666
22.4192 | -104.6657
22.4228 | -104.6671
22.4264 | -104.6703
22.43 | -104.6746
22.4337 | -104.68
22.4373 | -104.684
22.4409 | -104.6912
22.4446 | -104.6977
22.4482 | -104.7013
22.4518 | -104.7014
22.4554 | -104.7033
22.4591 | -104.694
22.4627 | -104.6767
22.4663 | -104.6484
22.47 | -104.6097
22.4736 | -104.5663
22.4772 | -104.5292
22.4808 | -104.5029
22.4845 | -104.4795
22.4881 | -104.4548
22.4917 | -104.4279
22.4954 | -104.3967
22.499 | -104.3641
22.5026 | -104.3281
22.5062 | -104.2634
22.5099 | -104.2195
22.5135 | -104.1738
22.5171 | -104.1263
22.5207 | -104.0768
22.5244 | -104.0236
22.528 | -103.9675
22.5316 | -103.9097
22.5353 | -103.8509
22.5389 | -103.791
22.5425 | -103.7293
22.5461 | -103.6658
22.5498 | -103.6012
22.5534 | -103.5375
22.557 | -103.4748
22.5607 | -103.4125
22.5643 | -103.3513
22.5679 | -103.2914
22.5715 | -103.2336
22.5752 | -103.1765
22.5788 | -103.1221
22.5824 | -103.0685
22.5861 | -103.0182
22.5897 | -102.9735
22.5933 | -102.9327
22.5969 | -102.8947
22.6006 | -102.8598
22.6042 | -102.8294
22.6078 | -102.8032
22.6115 | -102.782
22.6151 | -102.766
22.6187 | -102.7554
22.6223 | -102.7486
22.626 | -102.7459
22.6296 | -102.7489
22.6332 | -102.7563
22.6369 | -102.7686
22.6405 | -102.7866
22.6441 | -102.8093
22.6477 | -102.8353
22.6514 | -102.8665
22.655 | -102.9023
22.6586 | -102.9424
22.6623 | -102.9865
22.6659 | -103.0341
22.6695 | -103.0845
22.6731 | -103.1396
22.6768 | -103.1977
22.6804 | -103.2566
22.684 | -103.3175
22.6877 | -103.3803
22.6913 | -103.4443
22.6949 | -103.5106
22.6985 | -103.5779
22.7022 | -103.6449
22.7058 | -103.7134
22.7094 | -103.7841
22.713 | -103.8537
22.7167 | -103.9222
22.7203 | -103.9908
22.7239 | -104.058
22.7276 | -104.1224
22.7312 | -104.1865
22.7348 | -104.2486
22.7384 | -104.3077
22.7421 | -104.3638
22.7457 | -104.4169
22.7493 | -104.469
22.753 | -104.5193
22.7566 | -104.5679
22.7602 | -104.6117
22.7638 | -104.6525
22.7675 | -104.6922
22.7711 | -104.7325
22.7747 | -104.7703
22.7784 | -104.8137
22.782 | -104.8585
22.7856 | -104.9069
22.7892 | -104.957
22.7929 | -105.0021
22.7965 | -105.0293
22.8001 | -105.0382
22.8038 | -105.0428
22.8074 | -105.0456
22.811 | -105.0455
22.8146 | -105.0465
22.8183 | -105.0467
22.8219 | -105.0451
22.8255 | -105.0424
22.8292 | -105.0415
22.8328 | -105.0409
22.8364 | -105.0377
22.84 | -105.0336
22.8437 | -105.0298
22.8473 | -105.0253
22.8509 | -105.0219
22.8546 | -105.0177
22.8582 | -105.0102
22.8618 | -105.003
22.8654 | -104.998
22.8691 | -104.9935
22.8727 | -104.9897
22.8763 | -104.9824
22.8799 | -104.9733
22.8836 | -104.965
22.8872 | -104.9568
22.8908 | -104.9479
22.8945 | -104.9408
22.8981 | -104.9358
22.9017 | -104.9277
22.9053 | -104.9217
22.909 | -104.9132
22.9126 | -104.9048
22.9162 | -104.8988
22.9199 | -104.8954
22.9235 | -104.8927
22.9271 | -104.8874
22.9307 | -104.8805
22.9344 | -104.8769
22.938 | -104.8731
22.9416 | -104.868
22.9453 | -104.8632
22.9489 | -104.862
22.9525 | -104.8606
22.9561 | -104.8595
22.9598 | -104.8589
22.9634 | -104.8575
22.967 | -104.8574
22.9707 | -104.8563
22.9743 | -104.854
22.9779 | -104.8543
22.9815 | -104.8539
22.9852 | -104.8515
22.9888 | -104.852
22.9924 | -104.8541
22.9961 | -104.8544
22.9997 | -104.855
23.0033 | -104.8576
23.0069 | -104.86
23.0106 | -104.8609
23.0142 | -104.8625
23.0178 | -104.8635
23.0215 | -104.8633
23.0251 | -104.8651
23.0287 | -104.8638
23.0323 | -104.8635
23.036 | -104.8611
23.0396 | -104.86
23.0432 | -104.8602
23.0468 | -104.8607
23.0505 | -104.8611
23.0541 | -104.8597
23.0577 | -104.8547
23.0614 | -104.8507
23.065 | -104.8484
23.0686 | -104.8448
23.0722 | -104.8404
23.0759 | -104.8364
23.0795 | -104.8297
23.0831 | -104.8206
23.0868 | -104.8137
23.0904 | -104.807
23.094 | -104.8003
23.0976 | -104.7937
23.1013 | -104.7837
23.1049 | -104.7733
23.1085 | -104.7648
23.1122 | -104.7563
23.1158 | -104.746
23.1194 | -104.734
23.123 | -104.7209
23.1267 | -104.7082
23.1303 | -104.6946
23.1339 | -104.6814
23.1376 | -104.6686
23.1412 | -104.6565
23.1448 | -104.6448
23.1484 | -104.6332
23.1521 | -104.6219
23.1557 | -104.6085
23.1593 | -104.5952
23.163 | -104.5832
23.1666 | -104.5747
23.1702 | -104.5663
23.1738 | -104.558
23.1775 | -104.55
23.1811 | -104.5415
23.1847 | -104.5347
23.1884 | -104.5282
23.192 | -104.5215
23.1956 | -104.5146
23.1992 | -104.5097
23.2029 | -104.5073
23.2065 | -104.507
23.2101 | -104.506
23.2138 | -104.5059
23.2174 | -104.5087
23.221 | -104.5117
23.2246 | -104.5145
23.2283 | -104.5196
23.2319 | -104.5253
23.2355 | -104.5325
23.2391 | -104.5389
23.2428 | -104.5461
23.2464 | -104.5536
23.25 | -104.5616
23.2537 | -104.5716
23.2573 | -104.584
23.2609 | -104.5928
23.2645 | -104.6032
23.2682 | -104.6158
23.2718 | -104.629
23.2754 | -104.6408
23.2791 | -104.6528
23.2827 | -104.6659
23.2863 | -104.6799
23.2899 | -104.6931
23.2936 | -104.7064
23.2972 | -104.7203
23.3008 | -104.7349
23.3045 | -104.7461
23.3081 | -104.7577
23.3117 | -104.7712
23.3153 | -104.7812
23.319 | -104.79
23.3226 | -104.7968
23.3262 | -104.8047
23.3299 | -104.8114
23.3335 | -104.8175
23.3371 | -104.8243
23.3407 | -104.8289
23.3444 | -104.8312
23.348 | -104.8316
23.3516 | -104.8302
23.3553 | -104.8302
23.3589 | -104.8313
23.3625 | -104.8293
23.3661 | -104.8273
23.3698 | -104.8237
23.3734 | -104.818
23.377 | -104.8117
23.3807 | -104.8064
23.3843 | -104.8021
23.3879 | -104.7993
23.3915 | -104.7944
23.3952 | -104.7849
23.3988 | -104.7766
23.4024 | -104.7689
23.4061 | -104.7599
23.4097 | -104.7533
23.4133 | -104.7474
23.4169 | -104.7401
23.4206 | -104.7315
23.4242 | -104.7231
23.4278 | -104.7137
23.4314 | -104.7054
23.4351 | -104.6987
23.4387 | -104.6918
23.4423 | -104.6838
23.446 | -104.6772
23.4496 | -104.6723
23.4532 | -104.6698
23.4568 | -104.6676
23.4605 | -104.6636
23.4641 | -104.6585
23.4677 | -104.6559
23.4714 | -104.6538
23.475 | -104.6506
23.4786 | -104.6484
23.4822 | -104.6458
23.4859 | -104.6431
23.4895 | -104.6419
23.4931 | -104.6417
23.4968 | -104.6398
23.5004 | -104.6371
23.504 | -104.6369
23.5076 | -104.6362
23.5113 | -104.6345
23.5149 | -104.6346
23.5185 | -104.6347
23.5222 | -104.6348
23.5258 | -104.6366
23.5294 | -104.6381
23.533 | -104.6394
23.5367 | -104.6409
23.5403 | -104.6437
23.5439 | -104.6454
23.5476 | -104.6482
23.5512 | -104.651
23.5548 | -104.6602
23.5584 | -104.665
23.5621 | -104.6692
23.5657 | -104.6731
23.5693 | -104.6791
23.573 | -104.6861
23.5766 | -104.6933
23.5802 | -104.7026
23.5838 | -104.7113
23.5875 | -104.7209
23.5911 | -104.7293
23.5947 | -104.7397
23.5983 | -104.7496
23.602 | -104.7592
23.6056 | -104.7693
23.6092 | -104.7779
23.6129 | -104.7881
23.6165 | -104.8001
23.6201 | -104.8126
23.6237 | -104.825
23.6274 | -104.8363
23.631 | -104.8486
23.6346 | -104.8603
23.6383 | -104.8742
23.6419 | -104.8866
23.6455 | -104.8986
23.6491 | -104.9113
23.6528 | -104.9249
23.6564 | -104.9396
23.66 | -104.9541
23.6637 | -104.9677
23.6673 | -104.979
23.6709 | -104.989
23.6745 | -105.0002
23.6782 | -105.0103
23.6818 | -105.0221
23.6854 | -105.0335
23.6891 | -105.042
23.6927 | -105.0494
23.6963 | -105.0576
23.6999 | -105.0654
23.7036 | -105.0721
23.7072 | -105.0762
23.7108 | -105.0806
23.7145 | -105.0826
23.7181 | -105.0846
23.7217 | -105.0863
23.7253 | -105.088
23.729 | -105.0918
23.7326 | -105.0947
23.7362 | -105.0966
23.7399 | -105.0982
23.7435 | -105.0979
23.7471 | -105.0978
23.7507 | -105.0981
23.7544 | -105.097
23.758 | -105.0945
23.7616 | -105.0926
23.7653 | -105.0906
23.7689 | -105.0869
23.7725 | -105.0841
23.7761 | -105.083
23.7798 | -105.081
23.7834 | -105.0774
23.787 | -105.0738
23.7906 | -105.0713
23.7943 | -105.069
23.7979 | -105.067
23.8015 | -105.0642
23.8052 | -105.0616
23.8088 | -105.0601
23.8124 | -105.059
23.816 | -105.0588
23.8197 | -105.0613
23.8233 | -105.062
23.8269 | -105.0625
23.8306 | -105.0636
23.8342 | -105.0648
23.8378 | -105.065
23.8414 | -105.0673
23.8451 | -105.0715
23.8487 | -105.0743
23.8523 | -105.0774
23.856 | -105.0801
23.8596 | -105.0829
23.8632 | -105.0858
23.8668 | -105.0889
23.8705 | -105.0901
23.8741 | -105.0922
23.8777 | -105.0961
23.8814 | -105.098
23.885 | -105.1025
23.8886 | -105.1057
23.8922 | -105.1086
23.8959 | -105.1136
23.8995 | -105.119
23.9031 | -105.1236
23.9068 | -105.1284
23.9104 | -105.1327
23.914 | -105.1353
23.9176 | -105.139
23.9213 | -105.1447
23.9249 | -105.1498
23.9285 | -105.155
23.9322 | -105.1598
23.9358 | -105.1656
23.9394 | -105.1718
23.943 | -105.1776
23.9467 | -105.1845
23.9503 | -105.1903
23.9539 | -105.1952
23.9575 | -105.1983
23.9612 | -105.2022
23.9648 | -105.2066
23.9684 | -105.2093
23.9721 | -105.2131
23.9757 | -105.2166
23.9793 | -105.2189
23.9829 | -105.2204
23.9866 | -105.2213
23.9902 | -105.221
23.9938 | -105.219
23.9975 | -105.2182
24.0011 | -105.2175
24.0047 | -105.2158
24.0083 | -105.214
24.012 | -105.2115
24.0156 | -105.2097
24.0192 | -105.2052
24.0229 | -105.2012
24.0265 | -105.1979
24.0301 | -105.1936
24.0337 | -105.189
24.0374 | -105.1836
24.041 | -105.1781
24.0446 | -105.1734
24.0483 | -105.1668
24.0519 | -105.1616
24.0555 | -105.1553
24.0591 | -105.1479
24.0628 | -105.1421
24.0664 | -105.1343
24.07 | -105.1248
24.0737 | -105.1154
24.0773 | -105.1105
24.0809 | -105.1038
24.0845 | -105.0986
24.0882 | -105.0931
24.0918 | -105.0862
24.0954 | -105.0817
24.0991 | -105.0776
24.1027 | -105.0728
24.1063 | -105.0672
24.1099 | -105.0621
24.1136 | -105.0564
24.1172 | -105.0508
24.1208 | -105.0457
24.1244 | -105.044
24.1281 | -105.0406
24.1317 | -105.0369
24.1353 | -105.0326
24.139 | -105.0296
24.1426 | -105.0261
24.1462 | -105.0218
24.1498 | -105.0194
24.1535 | -105.0162
24.1571 | -105.0133
24.1607 | -105.0091
24.1644 | -105.003
24.168 | -105.0004
24.1716 | -104.9983
24.1752 | -104.9964
24.1789 | -104.9929
24.1825 | -104.9883
24.1861 | -104.9839
24.1898 | -104.9816
24.1934 | -104.9765
24.197 | -104.9727
24.2006 | -104.9697
24.2043 | -104.9665
24.2079 | -104.9638
24.2115 | -104.9613
24.2152 | -104.9587
24.2188 | -104.9559
24.2224 | -104.9541
24.226 | -104.9504
24.2297 | -104.9474
24.2333 | -104.9444
24.2369 | -104.9411
24.2406 | -104.9381
24.2442 | -104.9359
24.2478 | -104.9339
24.2514 | -104.9301
24.2551 | -104.9275
24.2587 | -104.9233
24.2623 | -104.9163
24.266 | -104.9071
24.2696 | -104.8939
24.2732 | -104.8742
24.2768 | -104.8437
24.2805 | -104.809
24.2841 | -104.7635
24.2877 | -104.724
24.2914 | -104.6931
24.295 | -104.6654
24.2986 | -104.6388
24.3022 | -104.6101
24.3059 | -104.5793
24.3095 | -104.5484
24.3131 | -104.5168
24.3167 | -104.4824
24.3204 | -104.448
24.324 | -104.4108
24.3276 | -104.37
24.3313 | -104.3291
24.3349 | -104.2879
24.3385 | -104.2426
24.3421 | -104.1957
24.3458 | -104.1488
24.3494 | -104.101
24.353 | -104.0517
24.3567 | -104.0008
24.3603 | -103.9493
24.3639 | -103.8974
24.3675 | -103.8463
24.3712 | -103.795
24.3748 | -103.7436
24.3784 | -103.6927
24.3821 | -103.6415
24.3857 | -103.5904
24.3893 | -103.5397
24.3929 | -103.4906
24.3966 | -103.4444
24.4002 | -103.4009
24.4038 | -103.3575
24.4075 | -103.3162
24.4111 | -103.2785
24.4147 | -103.2438
24.4183 | -103.2109
24.422 | -103.18
24.4256 | -103.152
24.4292 | -103.1271
24.4329 | -103.1067
24.4365 | -103.0902
24.4401 | -103.0756
24.4437 | -103.0651
24.4474 | -103.0585
24.451 | -103.0555
24.4546 | -103.0562
24.4583 | -103.0603
24.4619 | -103.0689
24.4655 | -103.0803
24.4691 | -103.0944
24.4728 | -103.1131
24.4764 | -103.1341
24.48 | -103.1603
24.4837 | -103.1894
24.4873 | -103.2209
24.4909 | -103.2558
24.4945 | -103.2935
24.4982 | -103.3335
24.5018 | -103.3761
24.5054 | -103.422
24.509 | -103.4703
24.5127 | -103.5178
24.5163 | -103.5672
24.5199 | -103.6179
24.5236 | -103.6699
24.5272 | -103.723
24.5308 | -103.7768
24.5344 | -103.8296
24.5381 | -103.8834
24.5417 | -103.9364
24.5453 | -103.9876
24.549 | -104.0378
24.5526 | -104.0863
24.5562 | -104.1308
24.5598 | -104.1699
24.5635 | -104.2025
24.5671 | -104.2213
24.5707 | -104.2332
24.5744 | -104.228
24.578 | -104.202
24.5816 | -104.1668
24.5852 | -104.1173
24.5889 | -104.072
24.5925 | -104.0363
24.5961 | -104.0043
24.5998 | -103.9069
24.6034 | -103.8582
24.607 | -103.805
24.6106 | -103.7494
24.6143 | -103.6884
24.6179 | -103.6254
24.6215 | -103.5591
24.6252 | -103.4933
24.6288 | -103.4237
24.6324 | -103.3537
24.636 | -103.2713
24.6397 | -103.176
24.6433 | -103.0735
24.6469 | -102.9681
24.6506 | -102.8617
24.6542 | -102.7553
24.6578 | -102.6481
24.6614 | -102.5411
24.6651 | -102.4352
24.6687 | -102.3306
24.6723 | -102.2278
24.6759 | -102.128
24.6796 | -102.0288
24.6832 | -101.9317
24.6868 | -101.8379
24.6905 | -101.7479
24.6941 | -101.6616
24.6977 | -101.5795
24.7013 | -101.5012
24.705 | -101.4275
24.7086 | -101.3585
24.7122 | -101.2936
24.7159 | -101.2333
24.7195 | -101.1792
24.7231 | -101.1303
24.7267 | -101.0867
24.7304 | -101.0498
24.734 | -101.0186
24.7376 | -100.9938
24.7413 | -100.9742
24.7449 | -100.96
24.7485 | -100.951
24.7521 | -100.9484
24.7558 | -100.9527
24.7594 | -100.9622
24.763 | -100.9783
24.7667 | -101.0002
24.7703 | -101.0284
24.7739 | -101.0625
24.7775 | -101.1026
24.7812 | -101.1483
24.7848 | -101.2002
24.7884 | -101.2581
24.7921 | -101.3211
24.7957 | -101.3892
24.7993 | -101.4623
24.8029 | -101.5403
24.8066 | -101.6224
24.8102 | -101.7084
24.8138 | -101.799
24.8175 | -101.8926
24.8211 | -101.9893
24.8247 | -102.0897
24.8283 | -102.1934
24.832 | -102.2992
24.8356 | -102.4074
24.8392 | -102.5175
24.8428 | -102.6287
24.8465 | -102.7398
24.8501 | -102.8479
24.8537 | -102.9573
24.8574 | -103.0689
24.861 | -103.1806
24.8646 | -103.2908
24.8682 | -103.4004
24.8719 | -103.508
24.8755 | -103.6119
24.8791 | -103.7121
24.8828 | -103.8085
24.8864 | -103.9034
24.89 | -103.9951
24.8936 | -104.0823
24.8973 | -104.166
24.9009 | -104.2463
24.9045 | -104.3244
24.9082 | -104.3995
24.9118 | -104.4731
24.9154 | -104.5508
24.919 | -104.6226
24.9227 | -104.698
24.9263 | -104.7807
24.9299 | -104.8558
24.9336 | -104.9289
24.9372 | -104.9769
24.9408 | -104.9955
24.9444 | -105.0007
24.9481 | -104.9984
24.9517 | -104.9984
24.9553 | -105.0004
24.959 | -104.9987
24.9626 | -104.9915
24.9662 | -104.9831
24.9698 | -104.9719
24.9735 | -104.9613
24.9771 | -104.9471
24.9807 | -104.9308
24.9844 | -104.911
24.988 | -104.8922
24.9916 | -104.8761
24.9952 | -104.8615
24.9989 | -104.847
25.0025 | -104.8331
25.0061 | -104.8185
25.0098 | -104.8028
25.0134 | -104.7855
25.017 | -104.7691
25.0206 | -104.7537
25.0243 | -104.7388
25.0279 | -104.7221
25.0315 | -104.7082
25.0351 | -104.6969
25.0388 | -104.6853
25.0424 | -104.6733
25.046 | -104.6611
25.0497 | -104.6489
25.0533 | -104.6372
25.0569 | -104.6263
25.0605 | -104.6162
25.0642 | -104.6055
25.0678 | -104.597
25.0714 | -104.5889
25.0751 | -104.5821
25.0787 | -104.576
25.0823 | -104.5693
25.0859 | -104.5639
25.0896 | -104.5591
25.0932 | -104.554
25.0968 | -104.5499
25.1005 | -104.5483
25.1041 | -104.5462
25.1077 | -104.543
25.1113 | -104.5432
25.115 | -104.5408
25.1186 | -104.5394
25.1222 | -104.5381
25.1259 | -104.5368
25.1295 | -104.5351
25.1331 | -104.5324
25.1367 | -104.5294
25.1404 | -104.5246
25.144 | -104.5191
25.1476 | -104.5148
25.1513 | -104.5104
25.1549 | -104.5077
25.1585 | -104.5035
25.1621 | -104.4992
25.1658 | -104.4959
25.1694 | -104.4928
25.173 | -104.4895
25.1767 | -104.4851
25.1803 | -104.4811
25.1839 | -104.4752
25.1875 | -104.4697
25.1912 | -104.4644
25.1948 | -104.4597
25.1984 | -104.4541
25.202 | -104.4519
25.2057 | -104.4524
25.2093 | -104.4489
25.2129 | -104.4421
25.2166 | -104.434
25.2202 | -104.426
25.2238 | -104.4171
25.2274 | -104.4079
25.2311 | -104.4016
25.2347 | -104.3965
25.2383 | -104.3924
25.242 | -104.3848
25.2456 | -104.3769
25.2492 | -104.3712
25.2528 | -104.3632
25.2565 | -104.3541
25.2601 | -104.3478
25.2637 | -104.3413
25.2674 | -104.3364
25.271 | -104.331
25.2746 | -104.3254
25.2782 | -104.3207
25.2819 | -104.3161
25.2855 | -104.3124
25.2891 | -104.307
25.2928 | -104.3028
25.2964 | -104.3014
25.3 | -104.2998
25.3036 | -104.2993
25.3073 | -104.2967
25.3109 | -104.2937
25.3145 | -104.2892
25.3182 | -104.2863
25.3218 | -104.2854
25.3254 | -104.2852
25.329 | -104.2849
25.3327 | -104.2848
25.3363 | -104.2849
25.3399 | -104.2878
25.3436 | -104.2902
25.3472 | -104.2914
25.3508 | -104.292
25.3544 | -104.2929
25.3581 | -104.2936
25.3617 | -104.2948
25.3653 | -104.2977
25.369 | -104.3007
25.3726 | -104.3036
25.3762 | -104.3071
25.3798 | -104.3113
25.3835 | -104.3157
25.3871 | -104.3201
25.3907 | -104.3246
25.3943 | -104.3298
25.398 | -104.3334
25.4016 | -104.3374
25.4052 | -104.3423
25.4089 | -104.3472
25.4125 | -104.352
25.4161 | -104.3563
25.4197 | -104.3611
25.4234 | -104.3639
25.427 | -104.3664
25.4306 | -104.3679
25.4343 | -104.3711
25.4379 | -104.3743
25.4415 | -104.377
25.4451 | -104.3777
25.4488 | -104.379
25.4524 | -104.3805
25.456 | -104.3793
25.4597 | -104.3776
25.4633 | -104.3775
25.4669 | -104.3751
25.4705 | -104.3729
25.4742 | -104.37
25.4778 | -104.3668
25.4814 | -104.364
25.4851 | -104.3605
25.4887 | -104.3557
25.4923 | -104.3511
25.4959 | -104.3458
25.4996 | -104.3417
25.5032 | -104.3359
25.5068 | -104.3288
25.5105 | -104.322
25.5141 | -104.3128
25.5177 | -104.3034
25.5213 | -104.2946
25.525 | -104.2845
25.5286 | -104.2727
25.5322 | -104.2624
25.5359 | -104.2527
25.5395 | -104.2403
25.5431 | -104.2296
25.5467 | -104.2174
25.5504 | -104.2047
25.554 | -104.1918
25.5576 | -104.1787
25.5613 | -104.1646
25.5649 | -104.152
25.5685 | -104.1393
25.5721 | -104.127
25.5758 | -104.1148
25.5794 | -104.1036
25.583 | -104.0937
25.5866 | -104.0844
25.5903 | -104.0724
25.5939 | -104.0585
25.5975 | -104.0454
25.6012 | -104.036
25.6048 | -104.0259
25.6084 | -104.015
25.612 | -104.0059
25.6157 | -103.9973
25.6193 | -103.9873
25.6229 | -103.9775
25.6266 | -103.9674
25.6302 | -103.9601
25.6338 | -103.9527
25.6374 | -103.9449
25.6411 | -103.937
25.6447 | -103.9281
25.6483 | -103.9271
25.652 | -103.9213
25.6556 | -103.9164
25.6592 | -103.9108
25.6628 | -103.905
25.6665 | -103.9014
25.6701 | -103.8996
25.6737 | -103.8976
25.6774 | -103.8974
25.681 | -103.8969
25.6846 | -103.8958
25.6882 | -103.8939
25.6919 | -103.8921
25.6955 | -103.8925
25.6991 | -103.8938
25.7028 | -103.8956
25.7064 | -103.8968
25.71 | -103.8997
25.7136 | -103.9031
25.7173 | -103.9079
25.7209 | -103.9131
25.7245 | -103.9174
25.7282 | -103.9225
25.7318 | -103.9288
25.7354 | -103.9353
25.739 | -103.9408
25.7427 | -103.9463
25.7463 | -103.9534
25.7499 | -103.961
25.7535 | -103.9679
25.7572 | -103.9747
25.7608 | -103.9842
25.7644 | -103.9926
25.7681 | -104
25.7717 | -104.0085
25.7753 | -104.0157
25.7789 | -104.0237
25.7826 | -104.0335
25.7862 | -104.043
25.7898 | -104.0517
25.7935 | -104.062
25.7971 | -104.0716
25.8007 | -104.0801
25.8043 | -104.0887
25.808 | -104.0982
25.8116 | -104.1072
25.8152 | -104.1152
25.8189 | -104.1248
25.8225 | -104.1342
25.8261 | -104.141
25.8297 | -104.1497
25.8334 | -104.1582
25.837 | -104.1674
25.8406 | -104.1773
25.8443 | -104.1846
25.8479 | -104.1925
25.8515 | -104.2011
25.8551 | -104.209
25.8588 | -104.216
25.8624 | -104.2224
25.866 | -104.2292
25.8697 | -104.2366
25.8733 | -104.2424
25.8769 | -104.2486
25.8805 | -104.2547
25.8842 | -104.2606
25.8878 | -104.2651
25.8914 | -104.2704
25.8951 | -104.2746
25.8987 | -104.2771
25.9023 | -104.2811
25.9059 | -104.2852
25.9096 | -104.2908
25.9132 | -104.2961
25.9168 | -104.2995
25.9204 | -104.3031
25.9241 | -104.3078
25.9277 | -104.3117
25.9313 | -104.315
25.935 | -104.3171
25.9386 | -104.3202
25.9422 | -104.3237
25.9458 | -104.3249
25.9495 | -104.3255
25.9531 | -104.3284
25.9567 | -104.3314
25.9604 | -104.3339
25.964 | -104.3349
25.9676 | -104.3373
25.9712 | -104.3388
25.9749 | -104.3413
25.9785 | -104.3448
25.9821 | -104.3463
25.9858 | -104.3484
25.9894 | -104.3514
25.993 | -104.354
25.9966 | -104.3591
26.0003 | -104.3628
26.0039 | -104.3673
26.0075 | -104.3714
26.0112 | -104.3734
26.0148 | -104.3748
26.0184 | -104.3782
26.022 | -104.3812
26.0257 | -104.383
26.0293 | -104.3864
26.0329 | -104.3897
26.0366 | -104.3931
26.0402 | -104.3966
26.0438 | -104.4004
26.0474 | -104.404
26.0511 | -104.4078
26.0547 | -104.4104
26.0583 | -104.4142
26.062 | -104.4188
26.0656 | -104.4232
26.0692 | -104.4281
26.0728 | -104.4327
26.0765 | -104.4347
26.0801 | -104.4389
26.0837 | -104.4429
26.0874 | -104.4472
26.091 | -104.4521
26.0946 | -104.4557
26.0982 | -104.4566
26.1019 | -104.4584
26.1055 | -104.4623
26.1091 | -104.466
26.1127 | -104.4709
26.1164 | -104.4764
26.12 | -104.482
26.1236 | -104.4877
26.1273 | -104.4925
26.1309 | -104.4983
26.1345 | -104.5032
26.1381 | -104.5068
26.1418 | -104.5133
26.1454 | -104.5191
26.149 | -104.5224
26.1527 | -104.5263
26.1563 | -104.5307
26.1599 | -104.5344
26.1635 | -104.5387
26.1672 | -104.5437
26.1708 | -104.5492
26.1744 | -104.5548
26.1781 | -104.56
26.1817 | -104.5659
26.1853 | -104.5717
26.1889 | -104.5755
26.1926 | -104.5801
26.1962 | -104.5866
26.1998 | -104.5915
26.2035 | -104.5964
26.2071 | -104.5997
26.2107 | -104.602
26.2143 | -104.6061
26.218 | -104.6089
26.2216 | -104.6112
26.2252 | -104.6153
26.2289 | -104.6189
26.2325 | -104.6218
26.2361 | -104.6249
26.2397 | -104.6267
26.2434 | -104.6283
26.247 | -104.6301
26.2506 | -104.6315
26.2543 | -104.6314
26.2579 | -104.633
26.2615 | -104.6341
26.2651 | -104.6346
26.2688 | -104.6356
26.2724 | -104.6364
26.276 | -104.6349
26.2797 | -104.631
26.2833 | -104.6257
26.2869 | -104.6214
26.2905 | -104.6186
26.2942 | -104.6148
26.2978 | -104.6104
26.3014 | -104.6074
26.305 | -104.6058
26.3087 | -104.6018
26.3123 | -104.5954
26.3159 | -104.5902
26.3196 | -104.5863
26.3232 | -104.5806
26.3268 | -104.5759
26.3304 | -104.5694
26.3341 | -104.5626
26.3377 | -104.5557
26.3413 | -104.5491
26.345 | -104.5411
26.3486 | -104.5325
26.3522 | -104.5259
26.3558 | -104.5192
26.3595 | -104.5112
26.3631 | -104.5025
26.3667 | -104.4938
26.3704 | -104.4856
26.374 | -104.4778
26.3776 | -104.471
26.3812 | -104.4645
26.3849 | -104.4579
26.3885 | -104.4506
26.3921 | -104.4421
26.3958 | -104.4348
26.3994 | -104.4269
26.403 | -104.4186
26.4066 | -104.4119
26.4103 | -104.4051
26.4139 | -104.3992
26.4175 | -104.3934
26.4212 | -104.3884
26.4248 | -104.3849
26.4284 | -104.3805
26.432 | -104.3756
26.4357 | -104.3697
26.4393 | -104.3642
26.4429 | -104.3616
26.4466 | -104.3618
26.4502 | -104.3599
26.4538 | -104.3572
26.4574 | -104.3551
26.4611 | -104.3544
26.4647 | -104.3534
26.4683 | -104.3555
26.4719 | -104.3574
26.4756 | -104.357
26.4792 | -104.3559
26.4828 | -104.356
26.4865 | -104.3566
26.4901 | -104.3577
26.4937 | -104.3605
26.4973 | -104.3638
26.501 | -104.3667
26.5046 | -104.3699
26.5082 | -104.3729
26.5119 | -104.3766
26.5155 | -104.3812
26.5191 | -104.3857
26.5227 | -104.3903
26.5264 | -104.395
26.53 | -104.4004
26.5336 | -104.4051
26.5373 | -104.4097
26.5409 | -104.4147
26.5445 | -104.4195
26.5481 | -104.4253
26.5518 | -104.4308
26.5554 | -104.4354
26.559 | -104.4389
26.5627 | -104.4431
26.5663 | -104.4476
26.5699 | -104.4521
26.5735 | -104.4567
26.5772 | -104.46
26.5808 | -104.4633
26.5844 | -104.4692
26.5881 | -104.4741
26.5917 | -104.4788
26.5953 | -104.4835
26.5989 | -104.4863
26.6026 | -104.4877
26.6062 | -104.4894
26.6098 | -104.4914
26.6135 | -104.4928
26.6171 | -104.4945
26.6207 | -104.4964
26.6243 | -104.4977
26.628 | -104.4995
26.6316 | -104.5
26.6352 | -104.5014
26.6389 | -104.5031
26.6425 | -104.5035
26.6461 | -104.5018
26.6497 | -104.5018
26.6534 | -104.5044
26.657 | -104.5054
26.6606 | -104.5055
26.6642 | -104.5041
26.6679 | -104.503
26.6715 | -104.503
26.6751 | -104.5015
26.6788 | -104.4992
26.6824 | -104.4991
26.686 | -104.4999
26.6896 | -104.5
26.6933 | -104.5004
26.6969 | -104.5003
26.7005 | -104.4995
26.7042 | -104.4969
26.7078 | -104.4955
26.7114 | -104.4945
26.715 | -104.4923
26.7187 | -104.4912
26.7223 | -104.4912
26.7259 | -104.4908
26.7296 | -104.491
26.7332 | -104.4913
26.7368 | -104.4906
26.7404 | -104.4888
26.7441 | -104.4874
26.7477 | -104.4868
26.7513 | -104.4852
26.755 | -104.4841
26.7586 | -104.4837
26.7622 | -104.4831
26.7658 | -104.4824
26.7695 | -104.4837
26.7731 | -104.4843
26.7767 | -104.4836
26.7804 | -104.4828
26.784 | -104.4818
26.7876 | -104.4821
26.7912 | -104.4798
26.7949 | -104.4778
26.7985 | -104.4772
26.8021 | -104.478
26.8058 | -104.4777
26.8094 | -104.4775
26.813 | -104.4781
26.8166 | -104.4773
26.8203 | -104.4764
26.8239 | -104.4757
26.8275 | -104.4759
26.8311 | -104.4763
26.8348 | -104.4777
26.8384 | -104.479
26.842 | -104.4778
26.8457 | -104.4759
26.8493 | -104.4746
26.8529 | -104.4744
26.8565 | -104.4764
26.8602 | -104.4759
26.8638 | -104.4734
26.8674 | -104.4703
26.8711 | -104.4668
26.8747 | -104.4628
26.8783 | -104.4618
26.8819 | -104.4615
26.8856 | -104.4616
26.8892 | -104.4589
26.8928 | -104.4561
26.8965 | -104.4533
26.9001 | -104.4496
26.9037 | -104.4469
26.9073 | -104.4423
26.911 | -104.438
26.9146 | -104.4339
26.9182 | -104.4286
26.9219 | -104.424
26.9255 | -104.4181
26.9291 | -104.4119
26.9327 | -104.4069
26.9364 | -104.4031
26.94 | -104.398
26.9436 | -104.393
26.9473 | -104.3881
26.9509 | -104.383
26.9545 | -104.3775
26.9581 | -104.3721
26.9618 | -104.3646
26.9654 | -104.3586
26.969 | -104.3533
26.9727 | -104.3471
26.9763 | -104.3403
26.9799 | -104.3362
26.9835 | -104.3321
26.9872 | -104.3279
26.9908 | -104.3246
26.9944 | -104.3206
26.998 | -104.3147
27.0017 | -104.3098
27.0053 | -104.3049
27.0089 | -104.3009
27.0126 | -104.2963
27.0162 | -104.2934
27.0198 | -104.2924
27.0234 | -104.2913
27.0271 | -104.289
27.0307 | -104.2871
27.0343 | -104.2878
27.038 | -104.2885
27.0416 | -104.2891
27.0452 | -104.2895
27.0488 | -104.2897
27.0525 | -104.2902
27.0561 | -104.2919
27.0597 | -104.2935
27.0634 | -104.2963
27.067 | -104.2986
27.0706 | -104.301
27.0742 | -104.3053
27.0779 | -104.3098
27.0815 | -104.3151
27.0851 | -104.3217
27.0888 | -104.3282
27.0924 | -104.334
27.096 | -104.3409
27.0996 | -104.3471
27.1033 | -104.3526
27.1069 | -104.3589
27.1105 | -104.3641
27.1142 | -104.3689
27.1178 | -104.3759
27.1214 | -104.3854
27.125 | -104.3941
27.1287 | -104.4031
27.1323 | -104.4124
27.1359 | -104.4202
27.1396 | -104.4276
27.1432 | -104.4357
27.1468 | -104.4441
27.1504 | -104.4516
27.1541 | -104.461
27.1577 | -104.4704
27.1613 | -104.4795
27.165 | -104.488
27.1686 | -104.497
27.1722 | -104.507
27.1758 | -104.5158
27.1795 | -104.5226
27.1831 | -104.5295
27.1867 | -104.5363
27.1903 | -104.5433
27.194 | -104.5503
27.1976 | -104.5578
27.2012 | -104.5623
27.2049 | -104.5669
27.2085 | -104.5736
27.2121 | -104.58
27.2157 | -104.5794
27.2194 | -104.5823
27.223 | -104.5853
27.2266 | -104.59
27.2303 | -104.5945
27.2339 | -104.5982
27.2375 | -104.6008
27.2411 | -104.6013
27.2448 | -104.6006
27.2484 | -104.5996
27.252 | -104.5998
27.2557 | -104.5997
27.2593 | -104.6007
27.2629 | -104.5987
27.2665 | -104.5955
27.2702 | -104.5927
27.2738 | -104.591
27.2774 | -104.5871
27.2811 | -104.583
27.2847 | -104.5782
27.2883 | -104.5729
27.2919 | -104.5665
27.2956 | -104.5583
27.2992 | -104.55
27.3028 | -104.5436
27.3065 | -104.5368
27.3101 | -104.5293
27.3137 | -104.5217
27.3173 | -104.513
27.321 | -104.5043
27.3246 | -104.4975
27.3282 | -104.4891
27.3319 | -104.4781
27.3355 | -104.4674
27.3391 | -104.4591
27.3427 | -104.4514
27.3464 | -104.4432
27.35 | -104.4335
27.3536 | -104.4239
27.3573 | -104.4144
27.3609 | -104.4053
27.3645 | -104.3943
27.3681 | -104.3825
27.3718 | -104.3717
27.3754 | -104.3615
27.379 | -104.3515
27.3826 | -104.3438
27.3863 | -104.3357
27.3899 | -104.3266
27.3935 | -104.318
27.3972 | -104.3085
27.4008 | -104.2987
27.4044 | -104.2907
27.408 | -104.2842
27.4117 | -104.2778
27.4153 | -104.271
27.4189 | -104.264
27.4226 | -104.2576
27.4262 | -104.2523
27.4298 | -104.2482
27.4334 | -104.2443
27.4371 | -104.2391
27.4407 | -104.2343
27.4443 | -104.2312
27.448 | -104.2282
27.4516 | -104.2267
27.4552 | -104.2244
27.4588 | -104.2226
27.4625 | -104.2214
27.4661 | -104.2227
27.4697 | -104.2221
27.4734 | -104.2224
27.477 | -104.2235
27.4806 | -104.2235
27.4842 | -104.2247
27.4879 | -104.2267
27.4915 | -104.2282
27.4951 | -104.2308
27.4988 | -104.2336
27.5024 | -104.2354
27.506 | -104.236
27.5096 | -104.2381
27.5133 | -104.2407
27.5169 | -104.2441
27.5205 | -104.2488
27.5242 | -104.2541
27.5278 | -104.2582
27.5314 | -104.2616
27.535 | -104.2656
27.5387 | -104.2671
27.5423 | -104.2701
27.5459 | -104.2724
27.5495 | -104.2757
27.5532 | -104.2784
27.5568 | -104.2804
27.5604 | -104.2827
27.5641 | -104.2873
27.5677 | -104.2915
27.5713 | -104.2955
27.5749 | -104.2979
27.5786 | -104.2995
27.5822 | -104.3013
27.5858 | -104.3059
27.5895 | -104.3088
27.5931 | -104.3104
27.5967 | -104.3123
27.6003 | -104.3147
27.604 | -104.318
27.6076 | -104.3194
27.6112 | -104.3203
27.6149 | -104.3223
27.6185 | -104.3244
27.6221 | -104.3259
27.6257 | -104.3274
27.6294 | -104.3297
27.633 | -104.3314
27.6366 | -104.3328
27.6403 | -104.3353
27.6439 | -104.3367
27.6475 | -104.3359
27.6511 | -104.3354
27.6548 | -104.3349
27.6584 | -104.3339
27.662 | -104.3335
27.6657 | -104.3345
27.6693 | -104.3354
27.6729 | -104.337
27.6765 | -104.3372
27.6802 | -104.3402
27.6838 | -104.342
27.6874 | -104.3445
27.6911 | -104.345
27.6947 | -104.3453
27.6983 | -104.3463
27.7019 | -104.3468
27.7056 | -104.3481
27.7092 | -104.3485
27.7128 | -104.3471
27.7164 | -104.3471
27.7201 | -104.3483
27.7237 | -104.3483
27.7273 | -104.3478
27.731 | -104.3456
27.7346 | -104.3446
27.7382 | -104.3455
27.7418 | -104.3444
27.7455 | -104.3422
27.7491 | -104.3422
27.7527 | -104.3428
27.7564 | -104.3441
27.76 | -104.3449
27.7636 | -104.3438
27.7672 | -104.3439
27.7709 | -104.3446
27.7745 | -104.3439
27.7781 | -104.3425
27.7818 | -104.3426
27.7854 | -104.3416
27.789 | -104.3429
27.7926 | -104.3437
27.7963 | -104.3428
27.7999 | -104.3421
27.8035 | -104.3414
27.8072 | -104.341
27.8108 | -104.3407
27.8144 | -104.3422
27.818 | -104.3433
27.8217 | -104.3431
27.8253 | -104.3434
27.8289 | -104.3457
27.8326 | -104.3485
27.8362 | -104.3508
27.8398 | -104.3522
27.8434 | -104.3536
27.8471 | -104.3571
27.8507 | -104.3604
27.8543 | -104.3618
27.858 | -104.3603
27.8616 | -104.3617
27.8652 | -104.3646
27.8688 | -104.3677
27.8725 | -104.3715
27.8761 | -104.3737
27.8797 | -104.3747
27.8834 | -104.3766
27.887 | -104.379
27.8906 | -104.3821
27.8942 | -104.385
27.8979 | -104.386
27.9015 | -104.3874
27.9051 | -104.39
27.9087 | -104.3915
27.9124 | -104.3903
27.916 | -104.3899
27.9196 | -104.3893
27.9233 | -104.387
27.9269 | -104.3755
27.9305 | -104.3637
27.9341 | -104.3561
27.9378 | -104.351
27.9414 | -104.3453
27.945 | -104.3397
27.9487 | -104.3344
27.9523 | -104.3281
27.9559 | -104.3169
27.9595 | -104.3062
27.9632 | -104.2965
27.9668 | -104.2868
27.9704 | -104.2771
27.9741 | -104.2655
27.9777 | -104.2527
27.9813 | -104.2397
27.9849 | -104.2253
27.9886 | -104.2105
27.9922 | -104.196
27.9958 | -104.1805
27.9995 | -104.162
28.0031 | -104.1447