## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 258 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 244
Hood | 02-92B/P | 204
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1012.8 | hPa |  | PRC PRO 01
Humidity | 55.2 | %/Hr
Internal temperature of calibration bench | 24.13 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1017 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.3 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.89 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.01 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3074 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 13/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 16/07/2018
Technical Manager | Anthony HUE | Approval |  | 16/07/2017
CEO | Robert GLEMOT | Approval |  | 16/07/2017
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4385 | 57.3592
0.012209 | -35.9874 | 38.3212
0.018313 | -34.9412 | 28.3592
0.024417 | -34.4921 | 22.0313
0.030521 | -34.3883 | 18.2641
0.036626 | -34.2485 | 15.1722
0.04273 | -34.1656 | 12.8611
0.048834 | -34.1155 | 11.2897
0.054938 | -34.0468 | 10.0937
0.061043 | -33.9721 | 8.9694
0.067147 | -33.9609 | 8.0976
0.073251 | -33.9614 | 7.3439
0.079355 | -33.9786 | 6.7876
0.08546 | -33.9958 | 6.334
0.091564 | -33.9919 | 5.9652
0.097668 | -33.958 | 5.7131
0.10377 | -33.9116 | 5.3647
0.10988 | -33.8815 | 5.0014
0.11598 | -33.8728 | 4.6406
0.12209 | -33.8893 | 4.3597
0.12819 | -33.9213 | 4.1249
0.13429 | -33.9522 | 3.885
0.1404 | -33.9686 | 3.6162
0.1465 | -33.9838 | 3.361
0.15261 | -33.9906 | 3.1236
0.15871 | -33.9872 | 2.9274
0.16482 | -33.9732 | 2.7684
0.17092 | -33.9494 | 2.7058
0.17702 | -33.9311 | 2.6085
0.18313 | -33.9143 | 2.5264
0.18923 | -33.9022 | 2.4217
0.19534 | -33.8953 | 2.3192
0.20144 | -33.8942 | 2.2067
0.20754 | -33.8968 | 2.132
0.21365 | -33.899 | 2.0569
0.21975 | -33.9031 | 1.9921
0.22586 | -33.9109 | 1.9273
0.23196 | -33.9192 | 1.856
0.23807 | -33.9275 | 1.78
0.24417 | -33.9295 | 1.6879
0.25027 | -33.9327 | 1.6045
0.25638 | -33.9326 | 1.5132
0.26248 | -33.9307 | 1.4264
0.26859 | -33.9283 | 1.343
0.27469 | -33.9276 | 1.2563
0.2808 | -33.9333 | 1.1922
0.2869 | -33.9377 | 1.1241
0.293 | -33.943 | 1.0671
0.29911 | -33.9492 | 1.0285
0.30521 | -33.955 | 0.99203
0.31132 | -33.959 | 0.97806
0.31742 | -33.9633 | 0.95574
0.32353 | -33.9667 | 0.93743
0.32963 | -33.9696 | 0.92203
0.33573 | -33.9715 | 0.90854
0.34184 | -33.9741 | 0.89152
0.34794 | -33.9774 | 0.85584
0.35405 | -33.9807 | 0.82539
0.36015 | -33.9843 | 0.7915
0.36626 | -33.9867 | 0.75391
0.37236 | -33.9891 | 0.70388
0.37846 | -33.9913 | 0.64648
0.38457 | -33.9903 | 0.58736
0.39067 | -33.9909 | 0.53695
0.39678 | -33.9893 | 0.48063
0.40288 | -33.9869 | 0.42738
0.40899 | -33.9831 | 0.37542
0.41509 | -33.9787 | 0.32339
0.42119 | -33.9737 | 0.27635
0.4273 | -33.97 | 0.22757
0.4334 | -33.9666 | 0.17811
0.43951 | -33.9635 | 0.12968
0.44561 | -33.962 | 0.084352
0.45172 | -33.9635 | 0.047091
0.45782 | -33.9632 | 0.0055834
0.46392 | -33.9644 | -0.032592
0.47003 | -33.966 | -0.067174
0.47613 | -33.9681 | -0.098865
0.48224 | -33.9703 | -0.12899
0.48834 | -33.9727 | -0.15035
0.49445 | -33.9747 | -0.17487
0.50055 | -33.9772 | -0.19719
0.50665 | -33.979 | -0.2154
0.51276 | -33.9801 | -0.23516
0.51886 | -33.9813 | -0.25249
0.52497 | -33.9804 | -0.25923
0.53107 | -33.9808 | -0.26966
0.53717 | -33.9812 | -0.28096
0.54328 | -33.981 | -0.28641
0.54938 | -33.981 | -0.29468
0.55549 | -33.9806 | -0.30382
0.56159 | -33.9805 | -0.31502
0.5677 | -33.9801 | -0.32558
0.5738 | -33.9796 | -0.33893
0.5799 | -33.9797 | -0.35112
0.58601 | -33.9801 | -0.36757
0.59211 | -33.9811 | -0.39003
0.59822 | -33.9817 | -0.40675
0.60432 | -33.9827 | -0.42838
0.61043 | -33.9836 | -0.44957
0.61653 | -33.9848 | -0.47114
0.62263 | -33.9858 | -0.49417
0.62874 | -33.9872 | -0.51452
0.63484 | -33.9883 | -0.53563
0.64095 | -33.9895 | -0.55671
0.64705 | -33.9906 | -0.57672
0.65316 | -33.9916 | -0.59625
0.65926 | -33.9926 | -0.61879
0.66536 | -33.9921 | -0.63723
0.67147 | -33.9923 | -0.65372
0.67757 | -33.9929 | -0.6718
0.68368 | -33.9927 | -0.69191
0.68978 | -33.9928 | -0.71001
0.69589 | -33.9925 | -0.72846
0.70199 | -33.9917 | -0.75034
0.70809 | -33.9912 | -0.77057
0.7142 | -33.9907 | -0.78993
0.7203 | -33.9899 | -0.80955
0.72641 | -33.9891 | -0.83039
0.73251 | -33.9889 | -0.85119
0.73862 | -33.9889 | -0.8754
0.74472 | -33.9891 | -0.8947
0.75082 | -33.9887 | -0.91575
0.75693 | -33.9888 | -0.93917
0.76303 | -33.9887 | -0.9616
0.76914 | -33.9892 | -0.98848
0.77524 | -33.9894 | -1.0133
0.78135 | -33.9898 | -1.0366
0.78745 | -33.9904 | -1.0613
0.79355 | -33.9909 | -1.0859
0.79966 | -33.9918 | -1.1115
0.80576 | -33.9929 | -1.1417
0.81187 | -33.9932 | -1.1711
0.81797 | -33.9938 | -1.2016
0.82408 | -33.9945 | -1.2287
0.83018 | -33.995 | -1.2597
0.83628 | -33.9958 | -1.29
0.84239 | -33.9957 | -1.318
0.84849 | -33.9957 | -1.3459
0.8546 | -33.9961 | -1.3734
0.8607 | -33.9961 | -1.4025
0.86681 | -33.9963 | -1.4306
0.87291 | -33.9958 | -1.4562
0.87901 | -33.9953 | -1.4785
0.88512 | -33.9952 | -1.5013
0.89122 | -33.9951 | -1.5245
0.89733 | -33.9948 | -1.5437
0.90343 | -33.9945 | -1.5648
0.90953 | -33.9947 | -1.5796
0.91564 | -33.9942 | -1.5954
0.92174 | -33.9939 | -1.6095
0.92785 | -33.9934 | -1.623
0.93395 | -33.9934 | -1.6346
0.94006 | -33.9934 | -1.6481
0.94616 | -33.9932 | -1.6528
0.95226 | -33.9934 | -1.6642
0.95837 | -33.9935 | -1.6758
0.96447 | -33.9939 | -1.6859
0.97058 | -33.9939 | -1.6957
0.97668 | -33.9938 | -1.7036
0.98279 | -33.9938 | -1.7073
0.98889 | -33.9941 | -1.7121
0.99499 | -33.9941 | -1.7206
1.0011 | -33.9944 | -1.7254
1.0072 | -33.9943 | -1.7292
1.0133 | -33.9943 | -1.7331
1.0194 | -33.9943 | -1.7386
1.0255 | -33.994 | -1.7455
1.0316 | -33.9934 | -1.7497
1.0377 | -33.993 | -1.7556
1.0438 | -33.9923 | -1.7602
1.0499 | -33.9908 | -1.7667
1.056 | -33.9897 | -1.7716
1.0621 | -33.9885 | -1.7775
1.0682 | -33.9881 | -1.7824
1.0743 | -33.987 | -1.7903
1.0805 | -33.9864 | -1.7972
1.0866 | -33.9848 | -1.808
1.0927 | -33.9835 | -1.8155
1.0988 | -33.9827 | -1.8261
1.1049 | -33.9813 | -1.8337
1.111 | -33.9797 | -1.8448
1.1171 | -33.9784 | -1.8542
1.1232 | -33.9771 | -1.8682
1.1293 | -33.976 | -1.8801
1.1354 | -33.975 | -1.8925
1.1415 | -33.9738 | -1.9065
1.1476 | -33.9725 | -1.9204
1.1537 | -33.9711 | -1.9336
1.1598 | -33.971 | -1.9518
1.1659 | -33.9705 | -1.9637
1.172 | -33.9698 | -1.977
1.1781 | -33.9691 | -1.9909
1.1842 | -33.9685 | -2.0063
1.1903 | -33.9692 | -2.0214
1.1964 | -33.9691 | -2.0358
1.2025 | -33.969 | -2.0509
1.2086 | -33.9695 | -2.0665
1.2147 | -33.9697 | -2.0814
1.2209 | -33.9704 | -2.0947
1.227 | -33.9722 | -2.1096
1.2331 | -33.9732 | -2.1238
1.2392 | -33.9741 | -2.1395
1.2453 | -33.9753 | -2.1514
1.2514 | -33.9766 | -2.1634
1.2575 | -33.9781 | -2.1773
1.2636 | -33.9804 | -2.1887
1.2697 | -33.9821 | -2.2003
1.2758 | -33.9836 | -2.2143
1.2819 | -33.9853 | -2.2276
1.288 | -33.987 | -2.2413
1.2941 | -33.989 | -2.2519
1.3002 | -33.991 | -2.2602
1.3063 | -33.9926 | -2.27
1.3124 | -33.9944 | -2.2799
1.3185 | -33.9963 | -2.2905
1.3246 | -33.9981 | -2.3012
1.3307 | -34.0002 | -2.3106
1.3368 | -34.002 | -2.3181
1.3429 | -34.0036 | -2.3268
1.349 | -34.0051 | -2.3352
1.3551 | -34.0068 | -2.346
1.3613 | -34.0083 | -2.3538
1.3674 | -34.0096 | -2.3622
1.3735 | -34.0108 | -2.3702
1.3796 | -34.012 | -2.3776
1.3857 | -34.013 | -2.3845
1.3918 | -34.0141 | -2.3942
1.3979 | -34.0154 | -2.4026
1.404 | -34.016 | -2.4126
1.4101 | -34.0169 | -2.4215
1.4162 | -34.0178 | -2.4301
1.4223 | -34.0185 | -2.4403
1.4284 | -34.0193 | -2.4487
1.4345 | -34.0199 | -2.4589
1.4406 | -34.0204 | -2.47
1.4467 | -34.0211 | -2.4794
1.4528 | -34.0217 | -2.4886
1.4589 | -34.0225 | -2.4991
1.465 | -34.0227 | -2.5068
1.4711 | -34.0237 | -2.5158
1.4772 | -34.0242 | -2.5279
1.4833 | -34.0248 | -2.5401
1.4894 | -34.0255 | -2.5515
1.4955 | -34.0262 | -2.563
1.5016 | -34.0269 | -2.575
1.5078 | -34.0273 | -2.5896
1.5139 | -34.028 | -2.6022
1.52 | -34.0285 | -2.6153
1.5261 | -34.0289 | -2.627
1.5322 | -34.0291 | -2.6396
1.5383 | -34.0294 | -2.6517
1.5444 | -34.0301 | -2.6651
1.5505 | -34.0305 | -2.6783
1.5566 | -34.031 | -2.6911
1.5627 | -34.0316 | -2.7045
1.5688 | -34.0324 | -2.7176
1.5749 | -34.0326 | -2.7299
1.581 | -34.033 | -2.7466
1.5871 | -34.0336 | -2.7621
1.5932 | -34.034 | -2.775
1.5993 | -34.0344 | -2.7918
1.6054 | -34.0347 | -2.8067
1.6115 | -34.0351 | -2.8236
1.6176 | -34.0355 | -2.8399
1.6237 | -34.0357 | -2.8543
1.6298 | -34.036 | -2.8697
1.6359 | -34.0364 | -2.8842
1.642 | -34.0369 | -2.8986
1.6482 | -34.0372 | -2.9156
1.6543 | -34.0377 | -2.9324
1.6604 | -34.0382 | -2.947
1.6665 | -34.0383 | -2.9619
1.6726 | -34.0386 | -2.9781
1.6787 | -34.0391 | -2.995
1.6848 | -34.0389 | -3.0114
1.6909 | -34.0393 | -3.0277
1.697 | -34.0395 | -3.0438
1.7031 | -34.0396 | -3.0595
1.7092 | -34.0396 | -3.0744
1.7153 | -34.0398 | -3.0905
1.7214 | -34.0394 | -3.1059
1.7275 | -34.0392 | -3.1217
1.7336 | -34.039 | -3.1378
1.7397 | -34.0391 | -3.1531
1.7458 | -34.0391 | -3.1697
1.7519 | -34.0389 | -3.1858
1.758 | -34.0388 | -3.2007
1.7641 | -34.0389 | -3.2151
1.7702 | -34.0387 | -3.2308
1.7763 | -34.0387 | -3.2445
1.7824 | -34.0386 | -3.2607
1.7885 | -34.0384 | -3.2747
1.7947 | -34.0385 | -3.2896
1.8008 | -34.0384 | -3.3024
1.8069 | -34.0383 | -3.316
1.813 | -34.0383 | -3.3302
1.8191 | -34.038 | -3.3445
1.8252 | -34.0378 | -3.3584
1.8313 | -34.0378 | -3.3728
1.8374 | -34.0378 | -3.3863
1.8435 | -34.0378 | -3.3984
1.8496 | -34.0375 | -3.4129
1.8557 | -34.0375 | -3.4276
1.8618 | -34.0374 | -3.4394
1.8679 | -34.0374 | -3.4524
1.874 | -34.0374 | -3.4656
1.8801 | -34.0373 | -3.4786
1.8862 | -34.0374 | -3.4914
1.8923 | -34.0373 | -3.5035
1.8984 | -34.037 | -3.5142
1.9045 | -34.0369 | -3.5253
1.9106 | -34.0366 | -3.5371
1.9167 | -34.0363 | -3.5484
1.9228 | -34.0363 | -3.5593
1.9289 | -34.036 | -3.5682
1.9351 | -34.036 | -3.5773
1.9412 | -34.0361 | -3.5877
1.9473 | -34.0362 | -3.5997
1.9534 | -34.0361 | -3.6108
1.9595 | -34.0358 | -3.6223
1.9656 | -34.0354 | -3.633
1.9717 | -34.0352 | -3.6435
1.9778 | -34.0352 | -3.6551
1.9839 | -34.0352 | -3.6665
1.99 | -34.0352 | -3.6764
1.9961 | -34.035 | -3.6868
2.0022 | -34.0349 | -3.6968
2.0083 | -34.0346 | -3.7067
2.0144 | -34.0345 | -3.7166
2.0205 | -34.0346 | -3.7284
2.0266 | -34.0347 | -3.7399
2.0327 | -34.0347 | -3.7511
2.0388 | -34.0349 | -3.7649
2.0449 | -34.035 | -3.7752
2.051 | -34.0351 | -3.7852
2.0571 | -34.0349 | -3.7966
2.0632 | -34.0348 | -3.8072
2.0693 | -34.0348 | -3.8183
2.0754 | -34.035 | -3.8301
2.0816 | -34.0352 | -3.8421
2.0877 | -34.0351 | -3.8531
2.0938 | -34.0354 | -3.8655
2.0999 | -34.0353 | -3.8771
2.106 | -34.0355 | -3.889
2.1121 | -34.0355 | -3.9013
2.1182 | -34.0357 | -3.9114
2.1243 | -34.0359 | -3.9219
2.1304 | -34.0359 | -3.9322
2.1365 | -34.0363 | -3.9437
2.1426 | -34.0364 | -3.957
2.1487 | -34.0369 | -3.9687
2.1548 | -34.0373 | -3.9817
2.1609 | -34.0378 | -3.9937
2.167 | -34.0383 | -4.0063
2.1731 | -34.0386 | -4.0182
2.1792 | -34.0393 | -4.0295
2.1853 | -34.0398 | -4.0423
2.1914 | -34.0402 | -4.0545
2.1975 | -34.0408 | -4.0667
2.2036 | -34.0412 | -4.0799
2.2097 | -34.0416 | -4.0919
2.2158 | -34.0426 | -4.1022
2.222 | -34.0431 | -4.1124
2.2281 | -34.0434 | -4.1232
2.2342 | -34.0439 | -4.1363
2.2403 | -34.0444 | -4.1464
2.2464 | -34.0451 | -4.1575
2.2525 | -34.0456 | -4.1683
2.2586 | -34.0464 | -4.1782
2.2647 | -34.0472 | -4.1876
2.2708 | -34.0478 | -4.1993
2.2769 | -34.0487 | -4.2108
2.283 | -34.0491 | -4.2203
2.2891 | -34.0497 | -4.2307
2.2952 | -34.0502 | -4.2406
2.3013 | -34.0508 | -4.2508
2.3074 | -34.0515 | -4.2605
2.3135 | -34.052 | -4.2719
2.3196 | -34.0526 | -4.2805
2.3257 | -34.0529 | -4.2911
2.3318 | -34.0532 | -4.3002
2.3379 | -34.0537 | -4.3116
2.344 | -34.054 | -4.3225
2.3501 | -34.0545 | -4.3333
2.3562 | -34.0548 | -4.3417
2.3623 | -34.055 | -4.3505
2.3685 | -34.0553 | -4.3596
2.3746 | -34.0555 | -4.3701
2.3807 | -34.056 | -4.3791
2.3868 | -34.0561 | -4.3882
2.3929 | -34.0565 | -4.3976
2.399 | -34.0568 | -4.4056
2.4051 | -34.0572 | -4.4152
2.4112 | -34.0575 | -4.4242
2.4173 | -34.0578 | -4.4332
2.4234 | -34.0578 | -4.4418
2.4295 | -34.0581 | -4.4505
2.4356 | -34.0582 | -4.4598
2.4417 | -34.0582 | -4.4699
2.4478 | -34.0583 | -4.4796
2.4539 | -34.0586 | -4.4875
2.46 | -34.0582 | -4.4973
2.4661 | -34.0583 | -4.5066
2.4722 | -34.0585 | -4.5154
2.4783 | -34.0586 | -4.5238
2.4844 | -34.0586 | -4.5325
2.4905 | -34.0585 | -4.5417
2.4966 | -34.0582 | -4.5507
2.5027 | -34.0582 | -4.5588
2.5089 | -34.058 | -4.5674
2.515 | -34.0578 | -4.5767
2.5211 | -34.0578 | -4.5851
2.5272 | -34.0574 | -4.5942
2.5333 | -34.0572 | -4.6037
2.5394 | -34.0572 | -4.6128
2.5455 | -34.057 | -4.6223
2.5516 | -34.0569 | -4.6311
2.5577 | -34.0568 | -4.6411
2.5638 | -34.0566 | -4.6514
2.5699 | -34.0564 | -4.6617
2.576 | -34.0561 | -4.6722
2.5821 | -34.0559 | -4.6822
2.5882 | -34.0557 | -4.6918
2.5943 | -34.0556 | -4.7016
2.6004 | -34.0554 | -4.7117
2.6065 | -34.0552 | -4.7202
2.6126 | -34.0551 | -4.7294
2.6187 | -34.055 | -4.7382
2.6248 | -34.0549 | -4.7473
2.6309 | -34.0547 | -4.7586
2.637 | -34.0547 | -4.7699
2.6431 | -34.0547 | -4.7794
2.6492 | -34.0545 | -4.7895
2.6554 | -34.0545 | -4.7998
2.6615 | -34.0543 | -4.8097
2.6676 | -34.0543 | -4.8203
2.6737 | -34.0544 | -4.8311
2.6798 | -34.0544 | -4.8411
2.6859 | -34.0544 | -4.8517
2.692 | -34.0543 | -4.8615
2.6981 | -34.0543 | -4.8727
2.7042 | -34.0543 | -4.8848
2.7103 | -34.0543 | -4.8949
2.7164 | -34.0545 | -4.9057
2.7225 | -34.0546 | -4.9171
2.7286 | -34.0547 | -4.9283
2.7347 | -34.0547 | -4.9384
2.7408 | -34.055 | -4.9477
2.7469 | -34.055 | -4.9591
2.753 | -34.0553 | -4.9698
2.7591 | -34.0554 | -4.9819
2.7652 | -34.0554 | -4.9935
2.7713 | -34.0554 | -5.0034
2.7774 | -34.0558 | -5.0149
2.7835 | -34.0559 | -5.0256
2.7896 | -34.0561 | -5.035
2.7958 | -34.0565 | -5.0454
2.8019 | -34.057 | -5.0553
2.808 | -34.0573 | -5.0672
2.8141 | -34.0577 | -5.0802
2.8202 | -34.0581 | -5.0924
2.8263 | -34.0585 | -5.1034
2.8324 | -34.0587 | -5.1159
2.8385 | -34.0592 | -5.1281
2.8446 | -34.0597 | -5.1398
2.8507 | -34.0601 | -5.152
2.8568 | -34.0606 | -5.1632
2.8629 | -34.061 | -5.1746
2.869 | -34.0614 | -5.186
2.8751 | -34.0618 | -5.1981
2.8812 | -34.0625 | -5.2114
2.8873 | -34.0628 | -5.2228
2.8934 | -34.0632 | -5.2347
2.8995 | -34.0635 | -5.2468
2.9056 | -34.0639 | -5.2576
2.9117 | -34.0643 | -5.2693
2.9178 | -34.065 | -5.282
2.9239 | -34.0654 | -5.2954
2.93 | -34.0659 | -5.3089
2.9361 | -34.0665 | -5.3204
2.9423 | -34.0671 | -5.3324
2.9484 | -34.0676 | -5.3459
2.9545 | -34.0683 | -5.3592
2.9606 | -34.0688 | -5.3726
2.9667 | -34.0693 | -5.3846
2.9728 | -34.0697 | -5.3969
2.9789 | -34.0701 | -5.4096
2.985 | -34.0706 | -5.4234
2.9911 | -34.0711 | -5.436
2.9972 | -34.0716 | -5.4488
3.0033 | -34.0722 | -5.4619
3.0094 | -34.0727 | -5.475
3.0155 | -34.0733 | -5.4876
3.0216 | -34.0738 | -5.501
3.0277 | -34.0742 | -5.5148
3.0338 | -34.0749 | -5.5284
3.0399 | -34.0754 | -5.5421
3.046 | -34.0758 | -5.5552
3.0521 | -34.0763 | -5.5677
3.0582 | -34.0766 | -5.5816
3.0643 | -34.0772 | -5.5947
3.0704 | -34.0774 | -5.6081
3.0765 | -34.0778 | -5.6204
3.0827 | -34.0781 | -5.6335
3.0888 | -34.0785 | -5.6464
3.0949 | -34.0788 | -5.6595
3.101 | -34.0793 | -5.6736
3.1071 | -34.0795 | -5.6877
3.1132 | -34.08 | -5.7018
3.1193 | -34.0804 | -5.7148
3.1254 | -34.0808 | -5.7271
3.1315 | -34.0812 | -5.7414
3.1376 | -34.0815 | -5.7543
3.1437 | -34.0818 | -5.7679
3.1498 | -34.0822 | -5.7821
3.1559 | -34.0825 | -5.7954
3.162 | -34.0825 | -5.8095
3.1681 | -34.0826 | -5.8237
3.1742 | -34.083 | -5.8374
3.1803 | -34.0834 | -5.8506
3.1864 | -34.0837 | -5.8629
3.1925 | -34.084 | -5.8766
3.1986 | -34.0841 | -5.8895
3.2047 | -34.0842 | -5.9035
3.2108 | -34.0845 | -5.9152
3.2169 | -34.0849 | -5.929
3.223 | -34.0849 | -5.9424
3.2292 | -34.0848 | -5.9563
3.2353 | -34.0848 | -5.9659
3.2414 | -34.0849 | -5.9789
3.2475 | -34.0849 | -5.99
3.2536 | -34.0852 | -6.0037
3.2597 | -34.0854 | -6.015
3.2658 | -34.0855 | -6.0258
3.2719 | -34.0857 | -6.0369
3.278 | -34.0857 | -6.0491
3.2841 | -34.0859 | -6.061
3.2902 | -34.086 | -6.0736
3.2963 | -34.0861 | -6.0853
3.3024 | -34.086 | -6.0955
3.3085 | -34.0862 | -6.1081
3.3146 | -34.0861 | -6.119
3.3207 | -34.0863 | -6.1312
3.3268 | -34.0864 | -6.1435
3.3329 | -34.0865 | -6.1533
3.339 | -34.0866 | -6.163
3.3451 | -34.0866 | -6.1745
3.3512 | -34.0868 | -6.184
3.3573 | -34.0869 | -6.1954
3.3634 | -34.087 | -6.2069
3.3696 | -34.0869 | -6.2188
3.3757 | -34.0869 | -6.2287
3.3818 | -34.0872 | -6.2404
3.3879 | -34.0872 | -6.2509
3.394 | -34.0871 | -6.2614
3.4001 | -34.0871 | -6.2723
3.4062 | -34.0871 | -6.2825
3.4123 | -34.0872 | -6.2924
3.4184 | -34.0872 | -6.3034
3.4245 | -34.0873 | -6.314
3.4306 | -34.0872 | -6.3238
3.4367 | -34.0874 | -6.3336
3.4428 | -34.0876 | -6.3428
3.4489 | -34.0876 | -6.3539
3.455 | -34.0875 | -6.3644
3.4611 | -34.0875 | -6.3756
3.4672 | -34.0875 | -6.3869
3.4733 | -34.0876 | -6.3971
3.4794 | -34.0876 | -6.4076
3.4855 | -34.0878 | -6.4173
3.4916 | -34.0879 | -6.427
3.4977 | -34.0878 | -6.438
3.5038 | -34.088 | -6.448
3.5099 | -34.088 | -6.4577
3.5161 | -34.0881 | -6.4672
3.5222 | -34.0882 | -6.4773
3.5283 | -34.0882 | -6.4874
3.5344 | -34.0882 | -6.4982
3.5405 | -34.0883 | -6.5081
3.5466 | -34.0884 | -6.518
3.5527 | -34.0885 | -6.5287
3.5588 | -34.0886 | -6.5385
3.5649 | -34.0886 | -6.5487
3.571 | -34.0887 | -6.5589
3.5771 | -34.0888 | -6.5697
3.5832 | -34.0889 | -6.5794
3.5893 | -34.0891 | -6.5893
3.5954 | -34.0892 | -6.5998
3.6015 | -34.0893 | -6.6102
3.6076 | -34.0895 | -6.6199
3.6137 | -34.0896 | -6.6298
3.6198 | -34.0899 | -6.6389
3.6259 | -34.09 | -6.6497
3.632 | -34.0902 | -6.6603
3.6381 | -34.0904 | -6.6702
3.6442 | -34.0906 | -6.6807
3.6503 | -34.0908 | -6.6898
3.6565 | -34.0911 | -6.7015
3.6626 | -34.0912 | -6.7123
3.6687 | -34.0915 | -6.7229
3.6748 | -34.0918 | -6.7328
3.6809 | -34.0921 | -6.7434
3.687 | -34.0923 | -6.7546
3.6931 | -34.0927 | -6.7646
3.6992 | -34.0929 | -6.7753
3.7053 | -34.0931 | -6.7854
3.7114 | -34.0932 | -6.7967
3.7175 | -34.0936 | -6.8066
3.7236 | -34.0938 | -6.8154
3.7297 | -34.0939 | -6.8262
3.7358 | -34.0941 | -6.8379
3.7419 | -34.0943 | -6.8496
3.748 | -34.0945 | -6.8603
3.7541 | -34.0946 | -6.8713
3.7602 | -34.095 | -6.8821
3.7663 | -34.0951 | -6.8923
3.7724 | -34.0955 | -6.9034
3.7785 | -34.0958 | -6.9138
3.7846 | -34.096 | -6.9245
3.7907 | -34.0964 | -6.9355
3.7969 | -34.0966 | -6.947
3.803 | -34.0969 | -6.9571
3.8091 | -34.0972 | -6.9674
3.8152 | -34.0975 | -6.9786
3.8213 | -34.0977 | -6.9903
3.8274 | -34.0981 | -7.0008
3.8335 | -34.0984 | -7.0124
3.8396 | -34.0989 | -7.0239
3.8457 | -34.0992 | -7.0342
3.8518 | -34.0997 | -7.0451
3.8579 | -34.1001 | -7.0557
3.864 | -34.1004 | -7.0665
3.8701 | -34.1008 | -7.0775
3.8762 | -34.1011 | -7.0887
3.8823 | -34.1014 | -7.0993
3.8884 | -34.1018 | -7.1097
3.8945 | -34.1023 | -7.1211
3.9006 | -34.1026 | -7.1326
3.9067 | -34.103 | -7.1445
3.9128 | -34.1034 | -7.1558
3.9189 | -34.1038 | -7.1665
3.925 | -34.104 | -7.1774
3.9311 | -34.1044 | -7.1874
3.9372 | -34.1048 | -7.199
3.9434 | -34.1051 | -7.2095
3.9495 | -34.1055 | -7.2211
3.9556 | -34.106 | -7.2323
3.9617 | -34.1063 | -7.2432
3.9678 | -34.1067 | -7.254
3.9739 | -34.1071 | -7.2652
3.98 | -34.1074 | -7.2771
3.9861 | -34.1077 | -7.2876
3.9922 | -34.1079 | -7.2992
3.9983 | -34.1083 | -7.3093
4.0044 | -34.1086 | -7.3191
4.0105 | -34.109 | -7.33
4.0166 | -34.1095 | -7.341
4.0227 | -34.1099 | -7.3514
4.0288 | -34.1101 | -7.3629
4.0349 | -34.1104 | -7.3728
4.041 | -34.1107 | -7.3838
4.0471 | -34.111 | -7.3941
4.0532 | -34.1115 | -7.4052
4.0593 | -34.1117 | -7.4168
4.0654 | -34.1121 | -7.4276
4.0715 | -34.1124 | -7.439
4.0776 | -34.1127 | -7.4494
4.0838 | -34.1131 | -7.4608
4.0899 | -34.1134 | -7.4721
4.096 | -34.1137 | -7.4826
4.1021 | -34.114 | -7.4934
4.1082 | -34.1144 | -7.5037
4.1143 | -34.1147 | -7.5137
4.1204 | -34.1152 | -7.5244
4.1265 | -34.1155 | -7.5349
4.1326 | -34.1159 | -7.5452
4.1387 | -34.1163 | -7.5564
4.1448 | -34.1167 | -7.5674
4.1509 | -34.1169 | -7.5782
4.157 | -34.1172 | -7.5894
4.1631 | -34.1175 | -7.6007
4.1692 | -34.1178 | -7.6114
4.1753 | -34.1182 | -7.6223
4.1814 | -34.1184 | -7.6322
4.1875 | -34.1187 | -7.6436
4.1936 | -34.119 | -7.6549
4.1997 | -34.1193 | -7.6651
4.2058 | -34.1196 | -7.6766
4.2119 | -34.1198 | -7.6864
4.218 | -34.1201 | -7.6965
4.2241 | -34.1205 | -7.7062
4.2303 | -34.1209 | -7.7174
4.2364 | -34.1212 | -7.7288
4.2425 | -34.1215 | -7.7394
4.2486 | -34.1217 | -7.7506
4.2547 | -34.1218 | -7.7598
4.2608 | -34.1221 | -7.7698
4.2669 | -34.1224 | -7.7802
4.273 | -34.1227 | -7.791
4.2791 | -34.1229 | -7.8018
4.2852 | -34.1231 | -7.8121
4.2913 | -34.1232 | -7.8214
4.2974 | -34.1235 | -7.8323
4.3035 | -34.1237 | -7.8432
4.3096 | -34.1238 | -7.8536
4.3157 | -34.1241 | -7.8636
4.3218 | -34.1243 | -7.8736
4.3279 | -34.1246 | -7.8841
4.334 | -34.1248 | -7.8939
4.3401 | -34.125 | -7.9039
4.3462 | -34.1253 | -7.9147
4.3523 | -34.1256 | -7.9244
4.3584 | -34.1257 | -7.9329
4.3645 | -34.1259 | -7.9426
4.3707 | -34.1261 | -7.952
4.3768 | -34.1263 | -7.9623
4.3829 | -34.1265 | -7.9727
4.389 | -34.1267 | -7.9824
4.3951 | -34.127 | -7.992
4.4012 | -34.1272 | -8.0019
4.4073 | -34.1274 | -8.012
4.4134 | -34.1276 | -8.0222
4.4195 | -34.1278 | -8.0319
4.4256 | -34.1281 | -8.0415
4.4317 | -34.1283 | -8.051
4.4378 | -34.1285 | -8.0606
4.4439 | -34.1287 | -8.0701
4.45 | -34.1288 | -8.0801
4.4561 | -34.1289 | -8.0899
4.4622 | -34.129 | -8.0996
4.4683 | -34.1292 | -8.1088
4.4744 | -34.1294 | -8.119
4.4805 | -34.1297 | -8.1281
4.4866 | -34.1299 | -8.1374
4.4927 | -34.13 | -8.1477
4.4988 | -34.1302 | -8.1571
4.5049 | -34.1304 | -8.1663
4.511 | -34.1307 | -8.1753
4.5172 | -34.131 | -8.1852
4.5233 | -34.1312 | -8.1958
4.5294 | -34.1314 | -8.2052
4.5355 | -34.1316 | -8.2146
4.5416 | -34.1319 | -8.2238
4.5477 | -34.132 | -8.2331
4.5538 | -34.1322 | -8.2418
4.5599 | -34.1325 | -8.2514
4.566 | -34.1327 | -8.2611
4.5721 | -34.1329 | -8.2704
4.5782 | -34.1332 | -8.2802
4.5843 | -34.1334 | -8.2899
4.5904 | -34.1337 | -8.2992
4.5965 | -34.134 | -8.308
4.6026 | -34.1343 | -8.3187
4.6087 | -34.1344 | -8.3277
4.6148 | -34.1347 | -8.3384
4.6209 | -34.135 | -8.3469
4.627 | -34.1354 | -8.357
4.6331 | -34.1354 | -8.3659
4.6392 | -34.1357 | -8.3752
4.6453 | -34.1359 | -8.3846
4.6514 | -34.1362 | -8.3942
4.6576 | -34.1362 | -8.4047
4.6637 | -34.1364 | -8.4124
4.6698 | -34.1367 | -8.421
4.6759 | -34.1368 | -8.4309
4.682 | -34.1369 | -8.4401
4.6881 | -34.1371 | -8.4489
4.6942 | -34.1374 | -8.4582
4.7003 | -34.1376 | -8.4677
4.7064 | -34.1378 | -8.477
4.7125 | -34.1381 | -8.4865
4.7186 | -34.1382 | -8.495
4.7247 | -34.1386 | -8.5046
4.7308 | -34.1389 | -8.514
4.7369 | -34.1391 | -8.5236
4.743 | -34.1393 | -8.534
4.7491 | -34.1394 | -8.5437
4.7552 | -34.1395 | -8.5529
4.7613 | -34.1397 | -8.5626
4.7674 | -34.14 | -8.5712
4.7735 | -34.1401 | -8.5814
4.7796 | -34.1403 | -8.5912
4.7857 | -34.1404 | -8.6009
4.7918 | -34.1407 | -8.6096
4.7979 | -34.141 | -8.6187
4.8041 | -34.1411 | -8.6281
4.8102 | -34.1415 | -8.6372
4.8163 | -34.1417 | -8.6467
4.8224 | -34.1418 | -8.6554
4.8285 | -34.1421 | -8.6645
4.8346 | -34.1423 | -8.6741
4.8407 | -34.1425 | -8.6835
4.8468 | -34.1428 | -8.6933
4.8529 | -34.143 | -8.7031
4.859 | -34.1432 | -8.7127
4.8651 | -34.1434 | -8.7221
4.8712 | -34.1436 | -8.7315
4.8773 | -34.1438 | -8.7411
4.8834 | -34.1441 | -8.7509
4.8895 | -34.1443 | -8.761
4.8956 | -34.1445 | -8.771
4.9017 | -34.1446 | -8.7805
4.9078 | -34.1448 | -8.7904
4.9139 | -34.1451 | -8.799
4.92 | -34.1453 | -8.8081
4.9261 | -34.1457 | -8.817
4.9322 | -34.146 | -8.8262
4.9383 | -34.1463 | -8.8353
4.9445 | -34.1466 | -8.8454
4.9506 | -34.1468 | -8.8544
4.9567 | -34.147 | -8.8637
4.9628 | -34.1473 | -8.8731
4.9689 | -34.1475 | -8.883
4.975 | -34.1477 | -8.8924
4.9811 | -34.1479 | -8.9023
4.9872 | -34.1482 | -8.9121
4.9933 | -34.1486 | -8.9218
4.9994 | -34.1489 | -8.9318
5.0055 | -34.1492 | -8.9414
5.0116 | -34.1495 | -8.9517
5.0177 | -34.1499 | -8.9619
5.0238 | -34.1502 | -8.9724
5.0299 | -34.1507 | -8.9834
5.036 | -34.1509 | -8.993
5.0421 | -34.1511 | -9.0028
5.0482 | -34.1514 | -9.0131
5.0543 | -34.1517 | -9.0225
5.0604 | -34.1521 | -9.0328
5.0665 | -34.1524 | -9.0425
5.0726 | -34.1527 | -9.053
5.0787 | -34.153 | -9.0633
5.0848 | -34.1533 | -9.0734
5.091 | -34.1536 | -9.0833
5.0971 | -34.1539 | -9.0934
5.1032 | -34.1543 | -9.1038
5.1093 | -34.1546 | -9.1138
5.1154 | -34.1548 | -9.1239
5.1215 | -34.155 | -9.1339
5.1276 | -34.1554 | -9.1438
5.1337 | -34.1558 | -9.1536
5.1398 | -34.1561 | -9.1634
5.1459 | -34.1564 | -9.1739
5.152 | -34.1567 | -9.185
5.1581 | -34.1569 | -9.1957
5.1642 | -34.1572 | -9.2057
5.1703 | -34.1576 | -9.2158
5.1764 | -34.1579 | -9.2258
5.1825 | -34.1582 | -9.2361
5.1886 | -34.1585 | -9.246
5.1947 | -34.1587 | -9.2568
5.2008 | -34.1592 | -9.2667
5.2069 | -34.1595 | -9.2779
5.213 | -34.1598 | -9.2876
5.2191 | -34.1602 | -9.298
5.2252 | -34.1605 | -9.3083
5.2314 | -34.1607 | -9.3191
5.2375 | -34.1611 | -9.3291
5.2436 | -34.1614 | -9.3393
5.2497 | -34.1617 | -9.3492
5.2558 | -34.1621 | -9.3603
5.2619 | -34.1624 | -9.3712
5.268 | -34.1627 | -9.3817
5.2741 | -34.1631 | -9.393
5.2802 | -34.1634 | -9.4038
5.2863 | -34.1638 | -9.4141
5.2924 | -34.1641 | -9.4253
5.2985 | -34.1645 | -9.4367
5.3046 | -34.165 | -9.448
5.3107 | -34.1653 | -9.4589
5.3168 | -34.1658 | -9.4694
5.3229 | -34.1661 | -9.4793
5.329 | -34.1664 | -9.4893
5.3351 | -34.1669 | -9.4995
5.3412 | -34.1673 | -9.5098
5.3473 | -34.1677 | -9.5207
5.3534 | -34.1681 | -9.5316
5.3595 | -34.1684 | -9.5424
5.3656 | -34.1687 | -9.5534
5.3717 | -34.169 | -9.564
5.3779 | -34.1694 | -9.575
5.384 | -34.1697 | -9.5858
5.3901 | -34.1701 | -9.5966
5.3962 | -34.1703 | -9.6074
5.4023 | -34.1708 | -9.6182
5.4084 | -34.1712 | -9.6287
5.4145 | -34.1715 | -9.6397
5.4206 | -34.1717 | -9.6511
5.4267 | -34.1721 | -9.6625
5.4328 | -34.1725 | -9.6727
5.4389 | -34.1729 | -9.683
5.445 | -34.1731 | -9.6944
5.4511 | -34.1735 | -9.7048
5.4572 | -34.1739 | -9.7161
5.4633 | -34.1743 | -9.7265
5.4694 | -34.1745 | -9.7371
5.4755 | -34.1749 | -9.748
5.4816 | -34.1752 | -9.7587
5.4877 | -34.1756 | -9.7694
5.4938 | -34.1759 | -9.7799
5.4999 | -34.1763 | -9.7906
5.506 | -34.1767 | -9.8014
5.5121 | -34.177 | -9.8126
5.5183 | -34.1773 | -9.8227
5.5244 | -34.1775 | -9.8336
5.5305 | -34.1779 | -9.8446
5.5366 | -34.1783 | -9.8551
5.5427 | -34.1786 | -9.8664
5.5488 | -34.1789 | -9.8774
5.5549 | -34.1792 | -9.8882
5.561 | -34.1795 | -9.8995
5.5671 | -34.1799 | -9.91
5.5732 | -34.1802 | -9.9209
5.5793 | -34.1806 | -9.9325
5.5854 | -34.181 | -9.9435
5.5915 | -34.1812 | -9.9545
5.5976 | -34.1815 | -9.9656
5.6037 | -34.1819 | -9.9764
5.6098 | -34.1822 | -9.987
5.6159 | -34.1825 | -9.998
5.622 | -34.1829 | -10.0085
5.6281 | -34.1832 | -10.0194
5.6342 | -34.1836 | -10.0303
5.6403 | -34.1838 | -10.0406
5.6464 | -34.1842 | -10.0515
5.6525 | -34.1845 | -10.062
5.6586 | -34.1849 | -10.0721
5.6648 | -34.1852 | -10.0834
5.6709 | -34.1856 | -10.0946
5.677 | -34.1859 | -10.1061
5.6831 | -34.1862 | -10.1165
5.6892 | -34.1865 | -10.1273
5.6953 | -34.1868 | -10.138
5.7014 | -34.1872 | -10.1494
5.7075 | -34.1875 | -10.1604
5.7136 | -34.1878 | -10.171
5.7197 | -34.1881 | -10.1817
5.7258 | -34.1884 | -10.1923
5.7319 | -34.1887 | -10.2031
5.738 | -34.1891 | -10.2138
5.7441 | -34.1894 | -10.2246
5.7502 | -34.1898 | -10.235
5.7563 | -34.1901 | -10.2455
5.7624 | -34.1904 | -10.257
5.7685 | -34.1907 | -10.2676
5.7746 | -34.1909 | -10.2782
5.7807 | -34.1912 | -10.2886
5.7868 | -34.1915 | -10.2993
5.7929 | -34.1918 | -10.31
5.799 | -34.1921 | -10.3206
5.8052 | -34.1924 | -10.3315
5.8113 | -34.1927 | -10.3422
5.8174 | -34.193 | -10.3531
5.8235 | -34.1933 | -10.3636
5.8296 | -34.1936 | -10.3749
5.8357 | -34.1938 | -10.3856
5.8418 | -34.1941 | -10.3966
5.8479 | -34.1943 | -10.4076
5.854 | -34.1946 | -10.4186
5.8601 | -34.1949 | -10.4291
5.8662 | -34.1952 | -10.4401
5.8723 | -34.1955 | -10.4502
5.8784 | -34.1958 | -10.4609
5.8845 | -34.1961 | -10.4717
5.8906 | -34.1964 | -10.4824
5.8967 | -34.1967 | -10.4927
5.9028 | -34.197 | -10.5034
5.9089 | -34.1973 | -10.5135
5.915 | -34.1976 | -10.5242
5.9211 | -34.1979 | -10.5349
5.9272 | -34.1982 | -10.5454
5.9333 | -34.1985 | -10.5554
5.9394 | -34.1988 | -10.5659
5.9455 | -34.1991 | -10.5761
5.9517 | -34.1995 | -10.5867
5.9578 | -34.1998 | -10.5977
5.9639 | -34.2001 | -10.608
5.97 | -34.2005 | -10.6187
5.9761 | -34.2008 | -10.629
5.9822 | -34.2011 | -10.6397
5.9883 | -34.2014 | -10.6501
5.9944 | -34.2017 | -10.6608
6.0005 | -34.2021 | -10.6716
6.0066 | -34.2024 | -10.6818
6.0127 | -34.2027 | -10.6917
6.0188 | -34.203 | -10.7026
6.0249 | -34.2033 | -10.7132
6.031 | -34.2036 | -10.7235
6.0371 | -34.204 | -10.7343
6.0432 | -34.2043 | -10.7448
6.0493 | -34.2047 | -10.7551
6.0554 | -34.205 | -10.7656
6.0615 | -34.2053 | -10.7766
6.0676 | -34.2056 | -10.7878
6.0737 | -34.206 | -10.7985
6.0798 | -34.2062 | -10.8089
6.0859 | -34.2066 | -10.8194
6.0921 | -34.2069 | -10.8296
6.0982 | -34.2073 | -10.8404
6.1043 | -34.2077 | -10.8507
6.1104 | -34.208 | -10.8618
6.1165 | -34.2083 | -10.8715
6.1226 | -34.2087 | -10.8815
6.1287 | -34.209 | -10.892
6.1348 | -34.2094 | -10.9026
6.1409 | -34.2097 | -10.9129
6.147 | -34.2101 | -10.9233
6.1531 | -34.2104 | -10.9337
6.1592 | -34.2108 | -10.9442
6.1653 | -34.2111 | -10.9544
6.1714 | -34.2114 | -10.9649
6.1775 | -34.2117 | -10.9757
6.1836 | -34.212 | -10.9859
6.1897 | -34.2124 | -10.9964
6.1958 | -34.2126 | -11.0071
6.2019 | -34.213 | -11.0178
6.208 | -34.2134 | -11.0287
6.2141 | -34.2137 | -11.0393
6.2202 | -34.2141 | -11.0503
6.2263 | -34.2144 | -11.0606
6.2325 | -34.2148 | -11.0705
6.2386 | -34.2152 | -11.0807
6.2447 | -34.2155 | -11.0914
6.2508 | -34.2158 | -11.1017
6.2569 | -34.2161 | -11.1125
6.263 | -34.2165 | -11.1226
6.2691 | -34.2169 | -11.1329
6.2752 | -34.2172 | -11.1432
6.2813 | -34.2176 | -11.1537
6.2874 | -34.2179 | -11.1641
6.2935 | -34.2183 | -11.1743
6.2996 | -34.2186 | -11.1848
6.3057 | -34.219 | -11.1951
6.3118 | -34.2193 | -11.2055
6.3179 | -34.2196 | -11.2159
6.324 | -34.22 | -11.2265
6.3301 | -34.2204 | -11.2367
6.3362 | -34.2207 | -11.2472
6.3423 | -34.2212 | -11.2577
6.3484 | -34.2216 | -11.2683
6.3545 | -34.2219 | -11.2781
6.3606 | -34.2223 | -11.2882
6.3667 | -34.2227 | -11.2984
6.3728 | -34.2231 | -11.3088
6.379 | -34.2234 | -11.3192
6.3851 | -34.2237 | -11.3296
6.3912 | -34.2241 | -11.3397
6.3973 | -34.2245 | -11.3498
6.4034 | -34.2248 | -11.3601
6.4095 | -34.2252 | -11.3707
6.4156 | -34.2255 | -11.3807
6.4217 | -34.2259 | -11.3911
6.4278 | -34.2262 | -11.4013
6.4339 | -34.2266 | -11.4116
6.44 | -34.227 | -11.4216
6.4461 | -34.2273 | -11.4324
6.4522 | -34.2277 | -11.4421
6.4583 | -34.228 | -11.4526
6.4644 | -34.2283 | -11.4633
6.4705 | -34.2287 | -11.4737
6.4766 | -34.2291 | -11.4845
6.4827 | -34.2294 | -11.4952
6.4888 | -34.2297 | -11.5054
6.4949 | -34.2301 | -11.5153
6.501 | -34.2304 | -11.5259
6.5071 | -34.2308 | -11.5364
6.5132 | -34.2313 | -11.5466
6.5194 | -34.2316 | -11.5568
6.5255 | -34.232 | -11.5668
6.5316 | -34.2323 | -11.5773
6.5377 | -34.2327 | -11.5878
6.5438 | -34.233 | -11.5981
6.5499 | -34.2334 | -11.6082
6.556 | -34.2338 | -11.6183
6.5621 | -34.2341 | -11.6289
6.5682 | -34.2344 | -11.639
6.5743 | -34.2348 | -11.6501
6.5804 | -34.2351 | -11.6604
6.5865 | -34.2354 | -11.671
6.5926 | -34.2358 | -11.6813
6.5987 | -34.2362 | -11.6912
6.6048 | -34.2365 | -11.7022
6.6109 | -34.2369 | -11.7125
6.617 | -34.2372 | -11.7236
6.6231 | -34.2375 | -11.7342
6.6292 | -34.2379 | -11.7444
6.6353 | -34.2383 | -11.7557
6.6414 | -34.2387 | -11.7664
6.6475 | -34.2391 | -11.7769
6.6536 | -34.2394 | -11.7869
6.6597 | -34.2398 | -11.797
6.6659 | -34.2402 | -11.8071
6.672 | -34.2406 | -11.8175
6.6781 | -34.2409 | -11.8277
6.6842 | -34.2413 | -11.8382
6.6903 | -34.2417 | -11.8485
6.6964 | -34.2421 | -11.8589
6.7025 | -34.2424 | -11.8692
6.7086 | -34.2427 | -11.8793
6.7147 | -34.2431 | -11.8898
6.7208 | -34.2434 | -11.9007
6.7269 | -34.2438 | -11.9106
6.733 | -34.244 | -11.9202
6.7391 | -34.2445 | -11.9308
6.7452 | -34.2449 | -11.9405
6.7513 | -34.2453 | -11.9505
6.7574 | -34.2457 | -11.9611
6.7635 | -34.2461 | -11.9709
6.7696 | -34.2465 | -11.9813
6.7757 | -34.2469 | -11.9919
6.7818 | -34.2473 | -12.0023
6.7879 | -34.2476 | -12.0127
6.794 | -34.248 | -12.0228
6.8001 | -34.2483 | -12.0329
6.8063 | -34.2487 | -12.0434
6.8124 | -34.249 | -12.0538
6.8185 | -34.2494 | -12.0641
6.8246 | -34.2497 | -12.0744
6.8307 | -34.2501 | -12.0845
6.8368 | -34.2505 | -12.0951
6.8429 | -34.2509 | -12.1053
6.849 | -34.2512 | -12.1156
6.8551 | -34.2516 | -12.126
6.8612 | -34.252 | -12.1361
6.8673 | -34.2523 | -12.1462
6.8734 | -34.2527 | -12.1572
6.8795 | -34.2531 | -12.1677
6.8856 | -34.2535 | -12.1775
6.8917 | -34.2538 | -12.1873
6.8978 | -34.2542 | -12.1975
6.9039 | -34.2545 | -12.2076
6.91 | -34.2549 | -12.2179
6.9161 | -34.2553 | -12.228
6.9222 | -34.2556 | -12.2383
6.9283 | -34.256 | -12.2482
6.9344 | -34.2563 | -12.2583
6.9405 | -34.2567 | -12.2683
6.9466 | -34.257 | -12.2788
6.9528 | -34.2574 | -12.2891
6.9589 | -34.2578 | -12.2992
6.965 | -34.2581 | -12.3094
6.9711 | -34.2585 | -12.3197
6.9772 | -34.2589 | -12.3297
6.9833 | -34.2592 | -12.3401
6.9894 | -34.2596 | -12.3507
6.9955 | -34.2599 | -12.3609
7.0016 | -34.2603 | -12.3715
7.0077 | -34.2607 | -12.382
7.0138 | -34.2611 | -12.3926
7.0199 | -34.2614 | -12.4026
7.026 | -34.2617 | -12.4127
7.0321 | -34.262 | -12.423
7.0382 | -34.2623 | -12.4329
7.0443 | -34.2628 | -12.4434
7.0504 | -34.263 | -12.454
7.0565 | -34.2634 | -12.4644
7.0626 | -34.2638 | -12.4744
7.0687 | -34.2641 | -12.4842
7.0748 | -34.2645 | -12.4943
7.0809 | -34.2648 | -12.5046
7.087 | -34.2651 | -12.5149
7.0932 | -34.2655 | -12.525
7.0993 | -34.2658 | -12.5353
7.1054 | -34.2661 | -12.5453
7.1115 | -34.2664 | -12.5553
7.1176 | -34.2667 | -12.5653
7.1237 | -34.267 | -12.5755
7.1298 | -34.2673 | -12.5853
7.1359 | -34.2676 | -12.5951
7.142 | -34.2679 | -12.6053
7.1481 | -34.2683 | -12.6152
7.1542 | -34.2686 | -12.6252
7.1603 | -34.2689 | -12.6352
7.1664 | -34.2693 | -12.6454
7.1725 | -34.2697 | -12.6553
7.1786 | -34.27 | -12.6655
7.1847 | -34.2703 | -12.6755
7.1908 | -34.2707 | -12.6858
7.1969 | -34.271 | -12.6958
7.203 | -34.2713 | -12.7059
7.2091 | -34.2717 | -12.7158
7.2152 | -34.272 | -12.7259
7.2213 | -34.2723 | -12.7357
7.2274 | -34.2727 | -12.7458
7.2335 | -34.273 | -12.7559
7.2397 | -34.2733 | -12.7659
7.2458 | -34.2736 | -12.7756
7.2519 | -34.274 | -12.7856
7.258 | -34.2743 | -12.7958
7.2641 | -34.2746 | -12.8058
7.2702 | -34.275 | -12.8161
7.2763 | -34.2754 | -12.826
7.2824 | -34.2758 | -12.8359
7.2885 | -34.2761 | -12.8459
7.2946 | -34.2764 | -12.8561
7.3007 | -34.2767 | -12.8661
7.3068 | -34.277 | -12.8764
7.3129 | -34.2774 | -12.8862
7.319 | -34.2777 | -12.8963
7.3251 | -34.2781 | -12.906
7.3312 | -34.2784 | -12.9157
7.3373 | -34.2788 | -12.9254
7.3434 | -34.2792 | -12.9352
7.3495 | -34.2795 | -12.9447
7.3556 | -34.2798 | -12.9551
7.3617 | -34.2801 | -12.9655
7.3678 | -34.2805 | -12.9756
7.3739 | -34.2808 | -12.9851
7.3801 | -34.2811 | -12.995
7.3862 | -34.2815 | -13.0049
7.3923 | -34.2818 | -13.0148
7.3984 | -34.2821 | -13.0247
7.4045 | -34.2825 | -13.0345
7.4106 | -34.2828 | -13.0443
7.4167 | -34.2831 | -13.0544
7.4228 | -34.2835 | -13.0643
7.4289 | -34.2838 | -13.0739
7.435 | -34.2842 | -13.0836
7.4411 | -34.2845 | -13.0932
7.4472 | -34.2849 | -13.1033
7.4533 | -34.2853 | -13.1131
7.4594 | -34.2856 | -13.1234
7.4655 | -34.286 | -13.1332
7.4716 | -34.2862 | -13.1433
7.4777 | -34.2866 | -13.1527
7.4838 | -34.2869 | -13.1628
7.4899 | -34.2873 | -13.1727
7.496 | -34.2877 | -13.1826
7.5021 | -34.288 | -13.1924
7.5082 | -34.2883 | -13.2024
7.5143 | -34.2887 | -13.2121
7.5204 | -34.289 | -13.2217
7.5266 | -34.2893 | -13.2316
7.5327 | -34.2896 | -13.2414
7.5388 | -34.2899 | -13.2512
7.5449 | -34.2902 | -13.2613
7.551 | -34.2906 | -13.2714
7.5571 | -34.2909 | -13.2811
7.5632 | -34.2912 | -13.2909
7.5693 | -34.2916 | -13.3009
7.5754 | -34.292 | -13.3108
7.5815 | -34.2923 | -13.3203
7.5876 | -34.2926 | -13.3301
7.5937 | -34.2929 | -13.3398
7.5998 | -34.2932 | -13.3498
7.6059 | -34.2935 | -13.3596
7.612 | -34.2939 | -13.3694
7.6181 | -34.2942 | -13.3791
7.6242 | -34.2946 | -13.389
7.6303 | -34.295 | -13.3987
7.6364 | -34.2953 | -13.4088
7.6425 | -34.2956 | -13.4189
7.6486 | -34.2959 | -13.4288
7.6547 | -34.2963 | -13.4387
7.6608 | -34.2966 | -13.4485
7.667 | -34.2969 | -13.4578
7.6731 | -34.2972 | -13.4678
7.6792 | -34.2975 | -13.4778
7.6853 | -34.2978 | -13.4875
7.6914 | -34.2981 | -13.4975
7.6975 | -34.2984 | -13.5075
7.7036 | -34.2987 | -13.5175
7.7097 | -34.2991 | -13.5272
7.7158 | -34.2994 | -13.5369
7.7219 | -34.2997 | -13.5467
7.728 | -34.3 | -13.5566
7.7341 | -34.3003 | -13.5665
7.7402 | -34.3007 | -13.5763
7.7463 | -34.301 | -13.5861
7.7524 | -34.3012 | -13.5962
7.7585 | -34.3016 | -13.6062
7.7646 | -34.3019 | -13.6158
7.7707 | -34.3021 | -13.6254
7.7768 | -34.3024 | -13.6351
7.7829 | -34.3027 | -13.6451
7.789 | -34.303 | -13.6546
7.7951 | -34.3033 | -13.6648
7.8012 | -34.3036 | -13.675
7.8073 | -34.3039 | -13.6847
7.8135 | -34.3042 | -13.6947
7.8196 | -34.3045 | -13.7044
7.8257 | -34.3049 | -13.7147
7.8318 | -34.3052 | -13.7249
7.8379 | -34.3055 | -13.7343
7.844 | -34.3057 | -13.7437
7.8501 | -34.306 | -13.7538
7.8562 | -34.3062 | -13.7643
7.8623 | -34.3066 | -13.7741
7.8684 | -34.3069 | -13.7837
7.8745 | -34.3072 | -13.7934
7.8806 | -34.3075 | -13.8028
7.8867 | -34.3078 | -13.8124
7.8928 | -34.3081 | -13.8221
7.8989 | -34.3085 | -13.8315
7.905 | -34.3088 | -13.8417
7.9111 | -34.3091 | -13.8516
7.9172 | -34.3095 | -13.8611
7.9233 | -34.3097 | -13.871
7.9294 | -34.31 | -13.881
7.9355 | -34.3103 | -13.8903
7.9416 | -34.3106 | -13.9
7.9477 | -34.3109 | -13.9094
7.9539 | -34.3112 | -13.9189
7.96 | -34.3115 | -13.9289
7.9661 | -34.3118 | -13.9388
7.9722 | -34.3122 | -13.9487
7.9783 | -34.3124 | -13.9585
7.9844 | -34.3128 | -13.9687
7.9905 | -34.3131 | -13.9783
7.9966 | -34.3133 | -13.9881
8.0027 | -34.3136 | -13.9983
8.0088 | -34.3139 | -14.0077
8.0149 | -34.3142 | -14.0172
8.021 | -34.3145 | -14.027
8.0271 | -34.3148 | -14.0373
8.0332 | -34.3151 | -14.0471
8.0393 | -34.3155 | -14.0571
8.0454 | -34.3157 | -14.067
8.0515 | -34.316 | -14.0768
8.0576 | -34.3163 | -14.0869
8.0637 | -34.3166 | -14.0962
8.0698 | -34.3169 | -14.1056
8.0759 | -34.3172 | -14.1151
8.082 | -34.3175 | -14.125
8.0881 | -34.3178 | -14.1347
8.0942 | -34.3181 | -14.1449
8.1004 | -34.3184 | -14.1551
8.1065 | -34.3187 | -14.1648
8.1126 | -34.319 | -14.1745
8.1187 | -34.3193 | -14.1844
8.1248 | -34.3196 | -14.1944
8.1309 | -34.3199 | -14.2042
8.137 | -34.3202 | -14.2139
8.1431 | -34.3206 | -14.2238
8.1492 | -34.3208 | -14.2337
8.1553 | -34.3212 | -14.2435
8.1614 | -34.3214 | -14.2536
8.1675 | -34.3217 | -14.2633
8.1736 | -34.322 | -14.2734
8.1797 | -34.3223 | -14.2833
8.1858 | -34.3225 | -14.2931
8.1919 | -34.3228 | -14.3028
8.198 | -34.3231 | -14.3123
8.2041 | -34.3235 | -14.3224
8.2102 | -34.3238 | -14.3322
8.2163 | -34.3241 | -14.3419
8.2224 | -34.3244 | -14.3519
8.2285 | -34.3247 | -14.3616
8.2346 | -34.325 | -14.3713
8.2408 | -34.3253 | -14.3812
8.2469 | -34.3256 | -14.3909
8.253 | -34.3259 | -14.4006
8.2591 | -34.3262 | -14.4101
8.2652 | -34.3266 | -14.4198
8.2713 | -34.3269 | -14.4296
8.2774 | -34.3272 | -14.4395
8.2835 | -34.3276 | -14.4497
8.2896 | -34.3278 | -14.459
8.2957 | -34.3281 | -14.4689
8.3018 | -34.3284 | -14.4786
8.3079 | -34.3288 | -14.4884
8.314 | -34.3291 | -14.498
8.3201 | -34.3294 | -14.5077
8.3262 | -34.3297 | -14.5175
8.3323 | -34.33 | -14.5272
8.3384 | -34.3303 | -14.5374
8.3445 | -34.3306 | -14.5468
8.3506 | -34.3309 | -14.5568
8.3567 | -34.3312 | -14.5669
8.3628 | -34.3315 | -14.5766
8.3689 | -34.3318 | -14.5864
8.375 | -34.3321 | -14.5962
8.3812 | -34.3324 | -14.606
8.3873 | -34.3327 | -14.6159
8.3934 | -34.3331 | -14.6256
8.3995 | -34.3334 | -14.6355
8.4056 | -34.3337 | -14.6454
8.4117 | -34.334 | -14.6555
8.4178 | -34.3343 | -14.6649
8.4239 | -34.3346 | -14.6744
8.43 | -34.3349 | -14.6842
8.4361 | -34.3352 | -14.6944
8.4422 | -34.3355 | -14.7041
8.4483 | -34.3358 | -14.7139
8.4544 | -34.3361 | -14.7239
8.4605 | -34.3364 | -14.7336
8.4666 | -34.3367 | -14.7436
8.4727 | -34.3369 | -14.7533
8.4788 | -34.3373 | -14.7629
8.4849 | -34.3376 | -14.7731
8.491 | -34.3378 | -14.7827
8.4971 | -34.3382 | -14.7926
8.5032 | -34.3385 | -14.8024
8.5093 | -34.3388 | -14.8122
8.5154 | -34.3391 | -14.8218
8.5215 | -34.3394 | -14.8316
8.5277 | -34.3397 | -14.8415
8.5338 | -34.34 | -14.851
8.5399 | -34.3403 | -14.8611
8.546 | -34.3406 | -14.8706
8.5521 | -34.3409 | -14.8805
8.5582 | -34.3412 | -14.8904
8.5643 | -34.3415 | -14.9
8.5704 | -34.3418 | -14.9095
8.5765 | -34.3421 | -14.9192
8.5826 | -34.3425 | -14.929
8.5887 | -34.3428 | -14.9387
8.5948 | -34.343 | -14.9481
8.6009 | -34.3434 | -14.9578
8.607 | -34.3437 | -14.9678
8.6131 | -34.3439 | -14.9778
8.6192 | -34.3442 | -14.9873
8.6253 | -34.3446 | -14.9972
8.6314 | -34.3449 | -15.0069
8.6375 | -34.3452 | -15.0169
8.6436 | -34.3456 | -15.0267
8.6497 | -34.3459 | -15.0366
8.6558 | -34.3462 | -15.0461
8.6619 | -34.3464 | -15.0561
8.6681 | -34.3467 | -15.0658
8.6742 | -34.347 | -15.0758
8.6803 | -34.3474 | -15.0858
8.6864 | -34.3477 | -15.0959
8.6925 | -34.3481 | -15.1056
8.6986 | -34.3484 | -15.1157
8.7047 | -34.3487 | -15.1256
8.7108 | -34.349 | -15.1355
8.7169 | -34.3493 | -15.1451
8.723 | -34.3497 | -15.1549
8.7291 | -34.3499 | -15.1646
8.7352 | -34.3503 | -15.1746
8.7413 | -34.3506 | -15.1843
8.7474 | -34.3509 | -15.1938
8.7535 | -34.3513 | -15.2037
8.7596 | -34.3516 | -15.2141
8.7657 | -34.3519 | -15.2238
8.7718 | -34.3522 | -15.2334
8.7779 | -34.3526 | -15.2437
8.784 | -34.3528 | -15.2534
8.7901 | -34.3531 | -15.2636
8.7962 | -34.3535 | -15.2735
8.8023 | -34.3538 | -15.283
8.8084 | -34.3541 | -15.293
8.8146 | -34.3544 | -15.3027
8.8207 | -34.3548 | -15.3125
8.8268 | -34.3551 | -15.3222
8.8329 | -34.3554 | -15.3322
8.839 | -34.3558 | -15.342
8.8451 | -34.3561 | -15.3517
8.8512 | -34.3565 | -15.3615
8.8573 | -34.3568 | -15.3711
8.8634 | -34.3572 | -15.3809
8.8695 | -34.3575 | -15.3908
8.8756 | -34.3578 | -15.4004
8.8817 | -34.3581 | -15.4104
8.8878 | -34.3584 | -15.4204
8.8939 | -34.3587 | -15.43
8.9 | -34.359 | -15.4395
8.9061 | -34.3594 | -15.4493
8.9122 | -34.3597 | -15.459
8.9183 | -34.36 | -15.4687
8.9244 | -34.3603 | -15.4783
8.9305 | -34.3607 | -15.4881
8.9366 | -34.361 | -15.4977
8.9427 | -34.3614 | -15.5076
8.9488 | -34.3617 | -15.5174
8.955 | -34.3621 | -15.5271
8.9611 | -34.3624 | -15.5371
8.9672 | -34.3628 | -15.5472
8.9733 | -34.3631 | -15.5568
8.9794 | -34.3634 | -15.5668
8.9855 | -34.3638 | -15.5765
8.9916 | -34.3641 | -15.5861
8.9977 | -34.3644 | -15.5958
9.0038 | -34.3648 | -15.6058
9.0099 | -34.3652 | -15.6157
9.016 | -34.3655 | -15.6255
9.0221 | -34.3659 | -15.6355
9.0282 | -34.3662 | -15.6454
9.0343 | -34.3666 | -15.6555
9.0404 | -34.367 | -15.6652
9.0465 | -34.3673 | -15.6753
9.0526 | -34.3677 | -15.685
9.0587 | -34.3681 | -15.6948
9.0648 | -34.3684 | -15.7045
9.0709 | -34.3688 | -15.714
9.077 | -34.3691 | -15.7244
9.0831 | -34.3694 | -15.7342
9.0892 | -34.3697 | -15.744
9.0953 | -34.37 | -15.7539
9.1015 | -34.3704 | -15.7637
9.1076 | -34.3708 | -15.7736
9.1137 | -34.3712 | -15.7833
9.1198 | -34.3715 | -15.7927
9.1259 | -34.3719 | -15.8025
9.132 | -34.3723 | -15.8126
9.1381 | -34.3727 | -15.8223
9.1442 | -34.3731 | -15.8317
9.1503 | -34.3734 | -15.8416
9.1564 | -34.3737 | -15.8515
9.1625 | -34.374 | -15.8613
9.1686 | -34.3744 | -15.8711
9.1747 | -34.3748 | -15.881
9.1808 | -34.3752 | -15.8911
9.1869 | -34.3755 | -15.9011
9.193 | -34.3759 | -15.9111
9.1991 | -34.3762 | -15.9206
9.2052 | -34.3765 | -15.9303
9.2113 | -34.3769 | -15.9404
9.2174 | -34.3773 | -15.9508
9.2235 | -34.3777 | -15.9604
9.2296 | -34.3781 | -15.9703
9.2357 | -34.3784 | -15.9801
9.2419 | -34.3788 | -15.9899
9.248 | -34.3792 | -15.9999
9.2541 | -34.3796 | -16.0099
9.2602 | -34.3799 | -16.0199
9.2663 | -34.3803 | -16.0295
9.2724 | -34.3806 | -16.0394
9.2785 | -34.381 | -16.0491
9.2846 | -34.3814 | -16.059
9.2907 | -34.3817 | -16.0688
9.2968 | -34.382 | -16.0788
9.3029 | -34.3824 | -16.0887
9.309 | -34.3828 | -16.0985
9.3151 | -34.3831 | -16.1081
9.3212 | -34.3836 | -16.1181
9.3273 | -34.384 | -16.1278
9.3334 | -34.3844 | -16.1374
9.3395 | -34.3848 | -16.1471
9.3456 | -34.3852 | -16.1567
9.3517 | -34.3856 | -16.1662
9.3578 | -34.386 | -16.1758
9.3639 | -34.3864 | -16.1858
9.37 | -34.3867 | -16.1956
9.3761 | -34.3871 | -16.2053
9.3822 | -34.3875 | -16.2151
9.3884 | -34.3879 | -16.2248
9.3945 | -34.3883 | -16.2347
9.4006 | -34.3886 | -16.2446
9.4067 | -34.389 | -16.2545
9.4128 | -34.3894 | -16.2642
9.4189 | -34.3898 | -16.2742
9.425 | -34.3902 | -16.2839
9.4311 | -34.3905 | -16.2937
9.4372 | -34.3909 | -16.3035
9.4433 | -34.3913 | -16.3134
9.4494 | -34.3917 | -16.3231
9.4555 | -34.3921 | -16.3329
9.4616 | -34.3925 | -16.3426
9.4677 | -34.3929 | -16.3523
9.4738 | -34.3933 | -16.3624
9.4799 | -34.3936 | -16.3722
9.486 | -34.3939 | -16.3823
9.4921 | -34.3943 | -16.3925
9.4982 | -34.3947 | -16.4019
9.5043 | -34.3951 | -16.4114
9.5104 | -34.3955 | -16.421
9.5165 | -34.3959 | -16.4308
9.5226 | -34.3963 | -16.4408
9.5288 | -34.3967 | -16.4506
9.5349 | -34.3971 | -16.4604
9.541 | -34.3975 | -16.4701
9.5471 | -34.3979 | -16.48
9.5532 | -34.3983 | -16.4898
9.5593 | -34.3987 | -16.4995
9.5654 | -34.3991 | -16.5094
9.5715 | -34.3995 | -16.5194
9.5776 | -34.3999 | -16.5294
9.5837 | -34.4002 | -16.5389
9.5898 | -34.4006 | -16.5489
9.5959 | -34.4009 | -16.5586
9.602 | -34.4013 | -16.5684
9.6081 | -34.4017 | -16.5783
9.6142 | -34.4021 | -16.588
9.6203 | -34.4024 | -16.5979
9.6264 | -34.4028 | -16.6073
9.6325 | -34.4031 | -16.6173
9.6386 | -34.4035 | -16.627
9.6447 | -34.4039 | -16.6368
9.6508 | -34.4043 | -16.6466
9.6569 | -34.4047 | -16.6563
9.663 | -34.4051 | -16.6662
9.6691 | -34.4055 | -16.676
9.6753 | -34.406 | -16.6858
9.6814 | -34.4064 | -16.6954
9.6875 | -34.4068 | -16.7049
9.6936 | -34.4072 | -16.7145
9.6997 | -34.4076 | -16.7245
9.7058 | -34.408 | -16.7343
9.7119 | -34.4084 | -16.744
9.718 | -34.4087 | -16.7541
9.7241 | -34.4091 | -16.7638
9.7302 | -34.4095 | -16.7737
9.7363 | -34.4099 | -16.7832
9.7424 | -34.4103 | -16.7929
9.7485 | -34.4107 | -16.8025
9.7546 | -34.4111 | -16.812
9.7607 | -34.4115 | -16.822
9.7668 | -34.4119 | -16.8321
9.7729 | -34.4123 | -16.8416
9.779 | -34.4128 | -16.8514
9.7851 | -34.4131 | -16.8613
9.7912 | -34.4135 | -16.8716
9.7973 | -34.4139 | -16.8815
9.8034 | -34.4143 | -16.8914
9.8095 | -34.4147 | -16.9011
9.8157 | -34.4151 | -16.9108
9.8218 | -34.4154 | -16.9203
9.8279 | -34.4159 | -16.9302
9.834 | -34.4163 | -16.9398
9.8401 | -34.4167 | -16.9494
9.8462 | -34.4171 | -16.9589
9.8523 | -34.4175 | -16.9685
9.8584 | -34.4179 | -16.9785
9.8645 | -34.4183 | -16.9883
9.8706 | -34.4187 | -16.998
9.8767 | -34.419 | -17.0073
9.8828 | -34.4195 | -17.0167
9.8889 | -34.4199 | -17.0267
9.895 | -34.4202 | -17.0361
9.9011 | -34.4206 | -17.0453
9.9072 | -34.421 | -17.0547
9.9133 | -34.4214 | -17.0642
9.9194 | -34.4219 | -17.0737
9.9255 | -34.4222 | -17.0835
9.9316 | -34.4227 | -17.093
9.9377 | -34.4231 | -17.1025
9.9438 | -34.4235 | -17.1121
9.9499 | -34.4239 | -17.1219
9.956 | -34.4243 | -17.1314
9.9622 | -34.4247 | -17.141
9.9683 | -34.4251 | -17.1507
9.9744 | -34.4255 | -17.1603
9.9805 | -34.4259 | -17.17
9.9866 | -34.4263 | -17.1794
9.9927 | -34.4268 | -17.1892
9.9988 | -34.4272 | -17.1987
10.0049 | -34.4275 | -17.2083
10.011 | -34.428 | -17.2179
10.0171 | -34.4284 | -17.2275
10.0232 | -34.4288 | -17.2369
10.0293 | -34.4293 | -17.2463
10.0354 | -34.4297 | -17.2556
10.0415 | -34.4301 | -17.2652
10.0476 | -34.4305 | -17.2752
10.0537 | -34.4309 | -17.2851
10.0598 | -34.4313 | -17.2948
10.0659 | -34.4317 | -17.3044
10.072 | -34.4321 | -17.3139
10.0781 | -34.4325 | -17.3235
10.0842 | -34.4329 | -17.3332
10.0903 | -34.4334 | -17.3431
10.0964 | -34.4338 | -17.3524
10.1026 | -34.4343 | -17.362
10.1087 | -34.4347 | -17.3716
10.1148 | -34.4351 | -17.3813
10.1209 | -34.4355 | -17.3908
10.127 | -34.4359 | -17.4001
10.1331 | -34.4364 | -17.4096
10.1392 | -34.4368 | -17.4191
10.1453 | -34.4372 | -17.4287
10.1514 | -34.4376 | -17.438
10.1575 | -34.438 | -17.4475
10.1636 | -34.4384 | -17.4564
10.1697 | -34.4388 | -17.4659
10.1758 | -34.4393 | -17.4753
10.1819 | -34.4397 | -17.485
10.188 | -34.4401 | -17.4943
10.1941 | -34.4404 | -17.5039
10.2002 | -34.4408 | -17.5129
10.2063 | -34.4412 | -17.5227
10.2124 | -34.4416 | -17.5322
10.2185 | -34.442 | -17.5417
10.2246 | -34.4425 | -17.5512
10.2307 | -34.4428 | -17.5609
10.2368 | -34.4433 | -17.5695
10.2429 | -34.4437 | -17.5788
10.2491 | -34.4441 | -17.5881
10.2552 | -34.4445 | -17.598
10.2613 | -34.445 | -17.6077
10.2674 | -34.4453 | -17.6174
10.2735 | -34.4457 | -17.6271
10.2796 | -34.4461 | -17.6367
10.2857 | -34.4465 | -17.6462
10.2918 | -34.4469 | -17.6559
10.2979 | -34.4474 | -17.6655
10.304 | -34.4478 | -17.6753
10.3101 | -34.4482 | -17.6847
10.3162 | -34.4487 | -17.6943
10.3223 | -34.4491 | -17.7038
10.3284 | -34.4495 | -17.7132
10.3345 | -34.4499 | -17.7225
10.3406 | -34.4503 | -17.7318
10.3467 | -34.4507 | -17.7412
10.3528 | -34.4511 | -17.7504
10.3589 | -34.4515 | -17.7601
10.365 | -34.4519 | -17.7698
10.3711 | -34.4523 | -17.7795
10.3772 | -34.4526 | -17.789
10.3833 | -34.453 | -17.7985
10.3895 | -34.4535 | -17.8078
10.3956 | -34.4539 | -17.817
10.4017 | -34.4543 | -17.8264
10.4078 | -34.4547 | -17.8359
10.4139 | -34.4551 | -17.8453
10.42 | -34.4556 | -17.8548
10.4261 | -34.456 | -17.8642
10.4322 | -34.4564 | -17.8734
10.4383 | -34.4568 | -17.8829
10.4444 | -34.4572 | -17.8919
10.4505 | -34.4577 | -17.9013
10.4566 | -34.458 | -17.9107
10.4627 | -34.4584 | -17.9201
10.4688 | -34.4589 | -17.929
10.4749 | -34.4593 | -17.9383
10.481 | -34.4597 | -17.9478
10.4871 | -34.4601 | -17.9573
10.4932 | -34.4606 | -17.9666
10.4993 | -34.461 | -17.9759
10.5054 | -34.4614 | -17.9853
10.5115 | -34.4619 | -17.9948
10.5176 | -34.4623 | -18.0039
10.5237 | -34.4627 | -18.0132
10.5298 | -34.4631 | -18.0224
10.536 | -34.4636 | -18.032
10.5421 | -34.464 | -18.041
10.5482 | -34.4644 | -18.0507
10.5543 | -34.4648 | -18.06
10.5604 | -34.4653 | -18.069
10.5665 | -34.4657 | -18.0782
10.5726 | -34.4661 | -18.0874
10.5787 | -34.4665 | -18.0969
10.5848 | -34.467 | -18.1059
10.5909 | -34.4673 | -18.1154
10.597 | -34.4677 | -18.1247
10.6031 | -34.4682 | -18.1342
10.6092 | -34.4686 | -18.1439
10.6153 | -34.469 | -18.1534
10.6214 | -34.4694 | -18.1626
10.6275 | -34.4699 | -18.1718
10.6336 | -34.4703 | -18.1815
10.6397 | -34.4707 | -18.1908
10.6458 | -34.4711 | -18.1999
10.6519 | -34.4714 | -18.2091
10.658 | -34.4718 | -18.2182
10.6641 | -34.4723 | -18.2276
10.6702 | -34.4727 | -18.2371
10.6764 | -34.4732 | -18.2464
10.6825 | -34.4736 | -18.2557
10.6886 | -34.4741 | -18.2651
10.6947 | -34.4745 | -18.2742
10.7008 | -34.4749 | -18.2835
10.7069 | -34.4753 | -18.2929
10.713 | -34.4757 | -18.3021
10.7191 | -34.4762 | -18.3116
10.7252 | -34.4766 | -18.3205
10.7313 | -34.4771 | -18.3298
10.7374 | -34.4775 | -18.3389
10.7435 | -34.4779 | -18.3484
10.7496 | -34.4783 | -18.3575
10.7557 | -34.4787 | -18.3668
10.7618 | -34.4792 | -18.3761
10.7679 | -34.4796 | -18.3852
10.774 | -34.48 | -18.3945
10.7801 | -34.4804 | -18.4042
10.7862 | -34.4808 | -18.4134
10.7923 | -34.4813 | -18.4228
10.7984 | -34.4817 | -18.4319
10.8045 | -34.4821 | -18.4413
10.8106 | -34.4825 | -18.4505
10.8168 | -34.483 | -18.4598
10.8229 | -34.4834 | -18.4691
10.829 | -34.4838 | -18.4785
10.8351 | -34.4842 | -18.4877
10.8412 | -34.4846 | -18.4969
10.8473 | -34.485 | -18.5058
10.8534 | -34.4854 | -18.5153
10.8595 | -34.4859 | -18.5243
10.8656 | -34.4863 | -18.5334
10.8717 | -34.4867 | -18.5423
10.8778 | -34.4872 | -18.5514
10.8839 | -34.4876 | -18.5606
10.89 | -34.488 | -18.5696
10.8961 | -34.4884 | -18.5789
10.9022 | -34.4889 | -18.5881
10.9083 | -34.4893 | -18.5976
10.9144 | -34.4898 | -18.6067
10.9205 | -34.4902 | -18.6159
10.9266 | -34.4907 | -18.6254
10.9327 | -34.4912 | -18.6346
10.9388 | -34.4916 | -18.6439
10.9449 | -34.492 | -18.6532
10.951 | -34.4924 | -18.6623
10.9571 | -34.4928 | -18.6717
10.9633 | -34.4932 | -18.6812
10.9694 | -34.4937 | -18.6908
10.9755 | -34.4941 | -18.6999
10.9816 | -34.4946 | -18.7092
10.9877 | -34.495 | -18.7184
10.9938 | -34.4954 | -18.7278
10.9999 | -34.4959 | -18.7369
11.006 | -34.4963 | -18.746
11.0121 | -34.4967 | -18.7556
11.0182 | -34.4971 | -18.7645
11.0243 | -34.4976 | -18.7737
11.0304 | -34.4979 | -18.7828
11.0365 | -34.4984 | -18.7919
11.0426 | -34.4988 | -18.8009
11.0487 | -34.4993 | -18.8102
11.0548 | -34.4997 | -18.8194
11.0609 | -34.5002 | -18.8288
11.067 | -34.5006 | -18.8382
11.0731 | -34.501 | -18.8474
11.0792 | -34.5015 | -18.8565
11.0853 | -34.5019 | -18.8655
11.0914 | -34.5023 | -18.8746
11.0975 | -34.5028 | -18.884
11.1037 | -34.5032 | -18.893
11.1098 | -34.5036 | -18.9025
11.1159 | -34.5041 | -18.9117
11.122 | -34.5045 | -18.9211
11.1281 | -34.5049 | -18.93
11.1342 | -34.5053 | -18.939
11.1403 | -34.5057 | -18.9484
11.1464 | -34.5062 | -18.9576
11.1525 | -34.5066 | -18.9669
11.1586 | -34.507 | -18.9766
11.1647 | -34.5075 | -18.9855
11.1708 | -34.5079 | -18.9945
11.1769 | -34.5083 | -19.0032
11.183 | -34.5087 | -19.0123
11.1891 | -34.5092 | -19.0218
11.1952 | -34.5096 | -19.0312
11.2013 | -34.5101 | -19.0404
11.2074 | -34.5105 | -19.0492
11.2135 | -34.5108 | -19.0584
11.2196 | -34.5112 | -19.0676
11.2257 | -34.5116 | -19.0767
11.2318 | -34.512 | -19.0861
11.2379 | -34.5124 | -19.0951
11.244 | -34.5128 | -19.1042
11.2502 | -34.5132 | -19.1134
11.2563 | -34.5136 | -19.1225
11.2624 | -34.514 | -19.1316
11.2685 | -34.5144 | -19.1406
11.2746 | -34.5148 | -19.1495
11.2807 | -34.5152 | -19.1593
11.2868 | -34.5156 | -19.1688
11.2929 | -34.5161 | -19.1776
11.299 | -34.5165 | -19.1869
11.3051 | -34.5169 | -19.196
11.3112 | -34.5173 | -19.2052
11.3173 | -34.5177 | -19.2144
11.3234 | -34.5181 | -19.2237
11.3295 | -34.5185 | -19.2326
11.3356 | -34.519 | -19.2418
11.3417 | -34.5194 | -19.2511
11.3478 | -34.5198 | -19.2602
11.3539 | -34.5202 | -19.2692
11.36 | -34.5206 | -19.2782
11.3661 | -34.521 | -19.2874
11.3722 | -34.5215 | -19.2964
11.3783 | -34.5219 | -19.3055
11.3844 | -34.5222 | -19.3147
11.3906 | -34.5227 | -19.3242
11.3967 | -34.5231 | -19.3332
11.4028 | -34.5235 | -19.3425
11.4089 | -34.5239 | -19.3514
11.415 | -34.5243 | -19.3605
11.4211 | -34.5247 | -19.3694
11.4272 | -34.5251 | -19.3786
11.4333 | -34.5256 | -19.3878
11.4394 | -34.526 | -19.3974
11.4455 | -34.5264 | -19.4065
11.4516 | -34.5268 | -19.4155
11.4577 | -34.5272 | -19.4248
11.4638 | -34.5277 | -19.4334
11.4699 | -34.5281 | -19.4429
11.476 | -34.5285 | -19.4526
11.4821 | -34.5289 | -19.4615
11.4882 | -34.5294 | -19.4709
11.4943 | -34.5298 | -19.4798
11.5004 | -34.5302 | -19.4889
11.5065 | -34.5307 | -19.498
11.5126 | -34.5311 | -19.5068
11.5187 | -34.5316 | -19.5158
11.5248 | -34.532 | -19.525
11.5309 | -34.5324 | -19.5342
11.5371 | -34.5327 | -19.5433
11.5432 | -34.5331 | -19.5526
11.5493 | -34.5335 | -19.5617
11.5554 | -34.5339 | -19.5706
11.5615 | -34.5343 | -19.5799
11.5676 | -34.5348 | -19.5893
11.5737 | -34.5352 | -19.5981
11.5798 | -34.5356 | -19.6071
11.5859 | -34.5361 | -19.616
11.592 | -34.5365 | -19.6251
11.5981 | -34.5369 | -19.634
11.6042 | -34.5373 | -19.6433
11.6103 | -34.5377 | -19.6523
11.6164 | -34.5381 | -19.6612
11.6225 | -34.5385 | -19.6703
11.6286 | -34.539 | -19.6795
11.6347 | -34.5394 | -19.6889
11.6408 | -34.5398 | -19.6981
11.6469 | -34.5403 | -19.7073
11.653 | -34.5407 | -19.7162
11.6591 | -34.5411 | -19.7253
11.6652 | -34.5415 | -19.7345
11.6713 | -34.5419 | -19.7435
11.6775 | -34.5424 | -19.7524
11.6836 | -34.5428 | -19.7616
11.6897 | -34.5432 | -19.7705
11.6958 | -34.5436 | -19.7799
11.7019 | -34.5441 | -19.7891
11.708 | -34.5445 | -19.7984
11.7141 | -34.545 | -19.8075
11.7202 | -34.5453 | -19.8164
11.7263 | -34.5457 | -19.8256
11.7324 | -34.5462 | -19.8348
11.7385 | -34.5466 | -19.8442
11.7446 | -34.547 | -19.8534
11.7507 | -34.5474 | -19.8625
11.7568 | -34.5478 | -19.8717
11.7629 | -34.5482 | -19.8808
11.769 | -34.5487 | -19.8898
11.7751 | -34.5491 | -19.8987
11.7812 | -34.5496 | -19.9076
11.7873 | -34.55 | -19.917
11.7934 | -34.5504 | -19.9261
11.7995 | -34.5508 | -19.9349
11.8056 | -34.5512 | -19.9444
11.8117 | -34.5516 | -19.9535
11.8178 | -34.552 | -19.9626
11.824 | -34.5525 | -19.9715
11.8301 | -34.553 | -19.9802
11.8362 | -34.5534 | -19.9889
11.8423 | -34.5537 | -19.998
11.8484 | -34.5542 | -20.007
11.8545 | -34.5546 | -20.0164
11.8606 | -34.555 | -20.0255
11.8667 | -34.5555 | -20.0346
11.8728 | -34.5558 | -20.0436
11.8789 | -34.5563 | -20.053
11.885 | -34.5566 | -20.062
11.8911 | -34.5571 | -20.0714
11.8972 | -34.5575 | -20.0801
11.9033 | -34.5579 | -20.0887
11.9094 | -34.5583 | -20.0975
11.9155 | -34.5587 | -20.1066
11.9216 | -34.5591 | -20.1155
11.9277 | -34.5596 | -20.1244
11.9338 | -34.5599 | -20.1335
11.9399 | -34.5604 | -20.1428
11.946 | -34.5608 | -20.1519
11.9521 | -34.5612 | -20.1608
11.9582 | -34.5617 | -20.1696
11.9644 | -34.5621 | -20.1785
11.9705 | -34.5625 | -20.1877
11.9766 | -34.563 | -20.1966
11.9827 | -34.5634 | -20.2056
11.9888 | -34.5638 | -20.2144
11.9949 | -34.5642 | -20.2234
12.001 | -34.5646 | -20.2325
12.0071 | -34.5651 | -20.2412
12.0132 | -34.5655 | -20.2499
12.0193 | -34.5659 | -20.2591
12.0254 | -34.5663 | -20.2679
12.0315 | -34.5668 | -20.2769
12.0376 | -34.5672 | -20.2856
12.0437 | -34.5675 | -20.2944
12.0498 | -34.568 | -20.3036
12.0559 | -34.5684 | -20.3127
12.062 | -34.5688 | -20.3216
12.0681 | -34.5692 | -20.3305
12.0742 | -34.5697 | -20.3396
12.0803 | -34.5701 | -20.3485
12.0864 | -34.5705 | -20.3575
12.0925 | -34.571 | -20.3665
12.0986 | -34.5714 | -20.3752
12.1047 | -34.5718 | -20.3842
12.1109 | -34.5723 | -20.3933
12.117 | -34.5727 | -20.4021
12.1231 | -34.5732 | -20.4111
12.1292 | -34.5736 | -20.4202
12.1353 | -34.5741 | -20.4289
12.1414 | -34.5745 | -20.438
12.1475 | -34.5749 | -20.447
12.1536 | -34.5753 | -20.4561
12.1597 | -34.5757 | -20.4652
12.1658 | -34.5761 | -20.4741
12.1719 | -34.5766 | -20.4831
12.178 | -34.577 | -20.4922
12.1841 | -34.5774 | -20.5013
12.1902 | -34.5779 | -20.5103
12.1963 | -34.5783 | -20.5192
12.2024 | -34.5788 | -20.5281
12.2085 | -34.5792 | -20.5371
12.2146 | -34.5796 | -20.5459
12.2207 | -34.58 | -20.5549
12.2268 | -34.5804 | -20.5638
12.2329 | -34.5809 | -20.5727
12.239 | -34.5813 | -20.5814
12.2451 | -34.5817 | -20.59
12.2513 | -34.5822 | -20.5991
12.2574 | -34.5826 | -20.6083
12.2635 | -34.583 | -20.6168
12.2696 | -34.5834 | -20.6255
12.2757 | -34.5839 | -20.6342
12.2818 | -34.5843 | -20.6429
12.2879 | -34.5846 | -20.6516
12.294 | -34.5851 | -20.6604
12.3001 | -34.5855 | -20.6694
12.3062 | -34.586 | -20.6783
12.3123 | -34.5864 | -20.6873
12.3184 | -34.5869 | -20.696
12.3245 | -34.5873 | -20.7049
12.3306 | -34.5877 | -20.7136
12.3367 | -34.5882 | -20.7223
12.3428 | -34.5886 | -20.731
12.3489 | -34.589 | -20.7393
12.355 | -34.5895 | -20.7482
12.3611 | -34.5899 | -20.7571
12.3672 | -34.5903 | -20.7658
12.3733 | -34.5907 | -20.7751
12.3794 | -34.5911 | -20.7834
12.3855 | -34.5915 | -20.7921
12.3916 | -34.592 | -20.8011
12.3978 | -34.5924 | -20.8102
12.4039 | -34.5928 | -20.8191
12.41 | -34.5932 | -20.8281
12.4161 | -34.5937 | -20.8367
12.4222 | -34.5941 | -20.8454
12.4283 | -34.5946 | -20.8542
12.4344 | -34.595 | -20.863
12.4405 | -34.5954 | -20.8719
12.4466 | -34.5958 | -20.8806
12.4527 | -34.5963 | -20.8895
12.4588 | -34.5967 | -20.8983
12.4649 | -34.5971 | -20.9074
12.471 | -34.5976 | -20.9166
12.4771 | -34.598 | -20.9249
12.4832 | -34.5984 | -20.9336
12.4893 | -34.5988 | -20.9423
12.4954 | -34.5993 | -20.951
12.5015 | -34.5998 | -20.9595
12.5076 | -34.6002 | -20.9684
12.5137 | -34.6006 | -20.9776
12.5198 | -34.601 | -20.9861
12.5259 | -34.6015 | -20.9948
12.532 | -34.602 | -21.0037
12.5382 | -34.6024 | -21.0124
12.5443 | -34.6028 | -21.0212
12.5504 | -34.6032 | -21.0301
12.5565 | -34.6037 | -21.0387
12.5626 | -34.6041 | -21.0478
12.5687 | -34.6046 | -21.0567
12.5748 | -34.605 | -21.0654
12.5809 | -34.6054 | -21.0739
12.587 | -34.6058 | -21.083
12.5931 | -34.6063 | -21.0917
12.5992 | -34.6067 | -21.1008
12.6053 | -34.6071 | -21.1094
12.6114 | -34.6076 | -21.1182
12.6175 | -34.608 | -21.1272
12.6236 | -34.6084 | -21.1357
12.6297 | -34.6089 | -21.1443
12.6358 | -34.6094 | -21.1528
12.6419 | -34.6098 | -21.1618
12.648 | -34.6103 | -21.1708
12.6541 | -34.6107 | -21.1794
12.6602 | -34.6112 | -21.1877
12.6663 | -34.6116 | -21.1963
12.6724 | -34.612 | -21.2051
12.6785 | -34.6124 | -21.2141
12.6847 | -34.6129 | -21.2228
12.6908 | -34.6133 | -21.2315
12.6969 | -34.6138 | -21.2402
12.703 | -34.6142 | -21.249
12.7091 | -34.6147 | -21.2578
12.7152 | -34.6151 | -21.2666
12.7213 | -34.6155 | -21.2754
12.7274 | -34.616 | -21.2842
12.7335 | -34.6164 | -21.2923
12.7396 | -34.6168 | -21.3006
12.7457 | -34.6173 | -21.3094
12.7518 | -34.6177 | -21.3181
12.7579 | -34.6181 | -21.3272
12.764 | -34.6185 | -21.3359
12.7701 | -34.6189 | -21.3444
12.7762 | -34.6193 | -21.3531
12.7823 | -34.6198 | -21.3618
12.7884 | -34.6201 | -21.3705
12.7945 | -34.6206 | -21.3789
12.8006 | -34.621 | -21.3869
12.8067 | -34.6215 | -21.3958
12.8128 | -34.6219 | -21.4041
12.8189 | -34.6222 | -21.4126
12.8251 | -34.6226 | -21.4211
12.8312 | -34.623 | -21.4298
12.8373 | -34.6235 | -21.4385
12.8434 | -34.6239 | -21.447
12.8495 | -34.6243 | -21.4556
12.8556 | -34.6247 | -21.4642
12.8617 | -34.6251 | -21.4727
12.8678 | -34.6256 | -21.4814
12.8739 | -34.626 | -21.4896
12.88 | -34.6264 | -21.4984
12.8861 | -34.6268 | -21.5071
12.8922 | -34.6273 | -21.5157
12.8983 | -34.6277 | -21.524
12.9044 | -34.6282 | -21.5326
12.9105 | -34.6287 | -21.5411
12.9166 | -34.6291 | -21.55
12.9227 | -34.6295 | -21.5591
12.9288 | -34.63 | -21.5678
12.9349 | -34.6305 | -21.5767
12.941 | -34.6309 | -21.5855
12.9471 | -34.6314 | -21.5941
12.9532 | -34.6319 | -21.6028
12.9593 | -34.6323 | -21.6116
12.9654 | -34.6327 | -21.6204
12.9716 | -34.6331 | -21.6291
12.9777 | -34.6336 | -21.6373
12.9838 | -34.634 | -21.6458
12.9899 | -34.6344 | -21.6545
12.996 | -34.6348 | -21.6632
13.0021 | -34.6352 | -21.6719
13.0082 | -34.6356 | -21.6805
13.0143 | -34.636 | -21.6887
13.0204 | -34.6364 | -21.697
13.0265 | -34.6369 | -21.7053
13.0326 | -34.6373 | -21.7141
13.0387 | -34.6378 | -21.7229
13.0448 | -34.6382 | -21.7315
13.0509 | -34.6387 | -21.7401
13.057 | -34.6391 | -21.7483
13.0631 | -34.6396 | -21.7567
13.0692 | -34.6401 | -21.7651
13.0753 | -34.6405 | -21.7737
13.0814 | -34.6409 | -21.7821
13.0875 | -34.6414 | -21.7903
13.0936 | -34.6418 | -21.7987
13.0997 | -34.6422 | -21.8069
13.1058 | -34.6427 | -21.8148
13.112 | -34.643 | -21.8234
13.1181 | -34.6435 | -21.8318
13.1242 | -34.6439 | -21.8404
13.1303 | -34.6444 | -21.8492
13.1364 | -34.6448 | -21.8576
13.1425 | -34.6452 | -21.866
13.1486 | -34.6456 | -21.8744
13.1547 | -34.6461 | -21.8826
13.1608 | -34.6465 | -21.891
13.1669 | -34.647 | -21.8994
13.173 | -34.6474 | -21.9079
13.1791 | -34.6478 | -21.9164
13.1852 | -34.6482 | -21.9252
13.1913 | -34.6487 | -21.9337
13.1974 | -34.6491 | -21.9421
13.2035 | -34.6495 | -21.9506
13.2096 | -34.65 | -21.959
13.2157 | -34.6504 | -21.9675
13.2218 | -34.6509 | -21.976
13.2279 | -34.6514 | -21.9845
13.234 | -34.6518 | -21.993
13.2401 | -34.6522 | -22.0014
13.2462 | -34.6526 | -22.0099
13.2524 | -34.653 | -22.0183
13.2585 | -34.6535 | -22.0267
13.2646 | -34.6539 | -22.0353
13.2707 | -34.6544 | -22.0438
13.2768 | -34.6548 | -22.0527
13.2829 | -34.6552 | -22.0612
13.289 | -34.6556 | -22.0697
13.2951 | -34.6561 | -22.078
13.3012 | -34.6565 | -22.0866
13.3073 | -34.6569 | -22.0953
13.3134 | -34.6573 | -22.1037
13.3195 | -34.6577 | -22.1121
13.3256 | -34.6582 | -22.1207
13.3317 | -34.6587 | -22.1289
13.3378 | -34.6591 | -22.1375
13.3439 | -34.6595 | -22.146
13.35 | -34.6599 | -22.1545
13.3561 | -34.6603 | -22.163
13.3622 | -34.6607 | -22.1714
13.3683 | -34.6612 | -22.1797
13.3744 | -34.6616 | -22.188
13.3805 | -34.662 | -22.1965
13.3866 | -34.6624 | -22.2047
13.3927 | -34.6629 | -22.213
13.3989 | -34.6633 | -22.2213
13.405 | -34.6638 | -22.2296
13.4111 | -34.6642 | -22.238
13.4172 | -34.6647 | -22.2462
13.4233 | -34.665 | -22.2549
13.4294 | -34.6655 | -22.263
13.4355 | -34.6659 | -22.272
13.4416 | -34.6663 | -22.2805
13.4477 | -34.6667 | -22.289
13.4538 | -34.6671 | -22.2973
13.4599 | -34.6676 | -22.3058
13.466 | -34.668 | -22.3146
13.4721 | -34.6685 | -22.3231
13.4782 | -34.6689 | -22.3314
13.4843 | -34.6694 | -22.3398
13.4904 | -34.6698 | -22.3477
13.4965 | -34.6702 | -22.3565
13.5026 | -34.6707 | -22.3648
13.5087 | -34.6711 | -22.3728
13.5148 | -34.6715 | -22.3814
13.5209 | -34.672 | -22.3896
13.527 | -34.6724 | -22.3983
13.5331 | -34.6728 | -22.4066
13.5393 | -34.6732 | -22.415
13.5454 | -34.6736 | -22.4235
13.5515 | -34.6741 | -22.4319
13.5576 | -34.6745 | -22.4405
13.5637 | -34.675 | -22.4489
13.5698 | -34.6754 | -22.4575
13.5759 | -34.6758 | -22.4655
13.582 | -34.6762 | -22.474
13.5881 | -34.6766 | -22.4825
13.5942 | -34.677 | -22.491
13.6003 | -34.6775 | -22.4994
13.6064 | -34.6779 | -22.5079
13.6125 | -34.6784 | -22.5167
13.6186 | -34.6789 | -22.525
13.6247 | -34.6793 | -22.5334
13.6308 | -34.6797 | -22.5415
13.6369 | -34.6801 | -22.5499
13.643 | -34.6806 | -22.5582
13.6491 | -34.681 | -22.5667
13.6552 | -34.6815 | -22.5748
13.6613 | -34.6819 | -22.5834
13.6674 | -34.6823 | -22.5919
13.6735 | -34.6828 | -22.6004
13.6796 | -34.6832 | -22.6086
13.6858 | -34.6836 | -22.6171
13.6919 | -34.6841 | -22.6255
13.698 | -34.6845 | -22.634
13.7041 | -34.6849 | -22.6426
13.7102 | -34.6853 | -22.6512
13.7163 | -34.6858 | -22.6597
13.7224 | -34.6863 | -22.668
13.7285 | -34.6867 | -22.6765
13.7346 | -34.6871 | -22.6849
13.7407 | -34.6875 | -22.6933
13.7468 | -34.688 | -22.7018
13.7529 | -34.6884 | -22.7101
13.759 | -34.6888 | -22.7186
13.7651 | -34.6892 | -22.7272
13.7712 | -34.6896 | -22.7357
13.7773 | -34.6901 | -22.7442
13.7834 | -34.6905 | -22.7526
13.7895 | -34.691 | -22.7609
13.7956 | -34.6914 | -22.7695
13.8017 | -34.6919 | -22.7777
13.8078 | -34.6923 | -22.7862
13.8139 | -34.6928 | -22.7945
13.82 | -34.6932 | -22.8026
13.8262 | -34.6937 | -22.8114
13.8323 | -34.6941 | -22.82
13.8384 | -34.6946 | -22.8286
13.8445 | -34.695 | -22.8371
13.8506 | -34.6955 | -22.8456
13.8567 | -34.6959 | -22.8544
13.8628 | -34.6963 | -22.8629
13.8689 | -34.6967 | -22.8714
13.875 | -34.6972 | -22.8797
13.8811 | -34.6976 | -22.8881
13.8872 | -34.6981 | -22.8964
13.8933 | -34.6986 | -22.9047
13.8994 | -34.699 | -22.9129
13.9055 | -34.6994 | -22.9213
13.9116 | -34.6999 | -22.9294
13.9177 | -34.7003 | -22.9374
13.9238 | -34.7008 | -22.946
13.9299 | -34.7013 | -22.9544
13.936 | -34.7017 | -22.9627
13.9421 | -34.7021 | -22.9711
13.9482 | -34.7026 | -22.9795
13.9543 | -34.703 | -22.9877
13.9604 | -34.7035 | -22.9962
13.9665 | -34.704 | -23.0047
13.9727 | -34.7044 | -23.013
13.9788 | -34.7048 | -23.0216
13.9849 | -34.7052 | -23.03
13.991 | -34.7056 | -23.0388
13.9971 | -34.706 | -23.0474
14.0032 | -34.7065 | -23.056
14.0093 | -34.7069 | -23.0643
14.0154 | -34.7073 | -23.0729
14.0215 | -34.7077 | -23.0811
14.0276 | -34.7081 | -23.0893
14.0337 | -34.7086 | -23.0979
14.0398 | -34.7091 | -23.1061
14.0459 | -34.7095 | -23.1147
14.052 | -34.7099 | -23.1231
14.0581 | -34.7103 | -23.1311
14.0642 | -34.7108 | -23.1397
14.0703 | -34.7112 | -23.1481
14.0764 | -34.7116 | -23.1569
14.0825 | -34.7121 | -23.1654
14.0886 | -34.7125 | -23.1734
14.0947 | -34.7129 | -23.1815
14.1008 | -34.7133 | -23.1903
14.1069 | -34.7138 | -23.1986
14.1131 | -34.7142 | -23.2073
14.1192 | -34.7146 | -23.2154
14.1253 | -34.715 | -23.2235
14.1314 | -34.7155 | -23.2319
14.1375 | -34.716 | -23.2405
14.1436 | -34.7165 | -23.2489
14.1497 | -34.7169 | -23.257
14.1558 | -34.7173 | -23.2654
14.1619 | -34.7178 | -23.2738
14.168 | -34.7182 | -23.282
14.1741 | -34.7187 | -23.2902
14.1802 | -34.7191 | -23.2984
14.1863 | -34.7196 | -23.3068
14.1924 | -34.72 | -23.3154
14.1985 | -34.7205 | -23.3235
14.2046 | -34.7209 | -23.3317
14.2107 | -34.7214 | -23.3398
14.2168 | -34.7218 | -23.3485
14.2229 | -34.7222 | -23.3569
14.229 | -34.7227 | -23.3653
14.2351 | -34.7231 | -23.3734
14.2412 | -34.7236 | -23.3816
14.2473 | -34.7241 | -23.3899
14.2534 | -34.7245 | -23.3982
14.2596 | -34.725 | -23.4066
14.2657 | -34.7254 | -23.4151
14.2718 | -34.7258 | -23.4236
14.2779 | -34.7263 | -23.4317
14.284 | -34.7267 | -23.44
14.2901 | -34.7272 | -23.4481
14.2962 | -34.7277 | -23.4562
14.3023 | -34.7281 | -23.4646
14.3084 | -34.7286 | -23.4732
14.3145 | -34.7291 | -23.4815
14.3206 | -34.7295 | -23.4899
14.3267 | -34.7299 | -23.4982
14.3328 | -34.7304 | -23.5065
14.3389 | -34.7308 | -23.5146
14.345 | -34.7312 | -23.5231
14.3511 | -34.7316 | -23.5312
14.3572 | -34.7321 | -23.5398
14.3633 | -34.7325 | -23.5481
14.3694 | -34.7329 | -23.5564
14.3755 | -34.7333 | -23.5644
14.3816 | -34.7338 | -23.5726
14.3877 | -34.7343 | -23.5808
14.3938 | -34.7347 | -23.5893
14.4 | -34.7352 | -23.5974
14.4061 | -34.7356 | -23.6057
14.4122 | -34.736 | -23.6142
14.4183 | -34.7364 | -23.6225
14.4244 | -34.7368 | -23.6306
14.4305 | -34.7373 | -23.6391
14.4366 | -34.7377 | -23.6476
14.4427 | -34.7382 | -23.6561
14.4488 | -34.7386 | -23.6641
14.4549 | -34.7391 | -23.6724
14.461 | -34.7395 | -23.6808
14.4671 | -34.74 | -23.6891
14.4732 | -34.7404 | -23.6971
14.4793 | -34.7409 | -23.706
14.4854 | -34.7414 | -23.7139
14.4915 | -34.7418 | -23.722
14.4976 | -34.7423 | -23.7304
14.5037 | -34.7428 | -23.7388
14.5098 | -34.7432 | -23.7469
14.5159 | -34.7437 | -23.7554
14.522 | -34.7441 | -23.7638
14.5281 | -34.7446 | -23.7718
14.5342 | -34.745 | -23.7802
14.5403 | -34.7455 | -23.7886
14.5465 | -34.7459 | -23.7972
14.5526 | -34.7464 | -23.8058
14.5587 | -34.7468 | -23.8141
14.5648 | -34.7473 | -23.8222
14.5709 | -34.7477 | -23.8307
14.577 | -34.7482 | -23.839
14.5831 | -34.7486 | -23.8474
14.5892 | -34.7491 | -23.8553
14.5953 | -34.7496 | -23.8633
14.6014 | -34.75 | -23.8716
14.6075 | -34.7505 | -23.8796
14.6136 | -34.7509 | -23.8876
14.6197 | -34.7514 | -23.896
14.6258 | -34.7518 | -23.9043
14.6319 | -34.7523 | -23.9124
14.638 | -34.7528 | -23.9206
14.6441 | -34.7532 | -23.9284
14.6502 | -34.7536 | -23.9364
14.6563 | -34.7541 | -23.9446
14.6624 | -34.7545 | -23.9528
14.6685 | -34.755 | -23.9607
14.6746 | -34.7555 | -23.969
14.6807 | -34.7559 | -23.9772
14.6869 | -34.7563 | -23.9854
14.693 | -34.7568 | -23.9934
14.6991 | -34.7572 | -24.0015
14.7052 | -34.7577 | -24.0096
14.7113 | -34.7581 | -24.0181
14.7174 | -34.7586 | -24.0262
14.7235 | -34.7591 | -24.0342
14.7296 | -34.7595 | -24.0424
14.7357 | -34.76 | -24.0504
14.7418 | -34.7604 | -24.0583
14.7479 | -34.7609 | -24.0666
14.754 | -34.7614 | -24.0748
14.7601 | -34.7618 | -24.083
14.7662 | -34.7623 | -24.0913
14.7723 | -34.7627 | -24.0993
14.7784 | -34.7632 | -24.1076
14.7845 | -34.7636 | -24.1158
14.7906 | -34.7641 | -24.1241
14.7967 | -34.7645 | -24.1322
14.8028 | -34.765 | -24.1402
14.8089 | -34.7654 | -24.1483
14.815 | -34.7658 | -24.1562
14.8211 | -34.7663 | -24.1643
14.8272 | -34.7667 | -24.1726
14.8334 | -34.7672 | -24.1806
14.8395 | -34.7677 | -24.1888
14.8456 | -34.7681 | -24.1968
14.8517 | -34.7686 | -24.2051
14.8578 | -34.769 | -24.2132
14.8639 | -34.7695 | -24.2215
14.87 | -34.77 | -24.2297
14.8761 | -34.7705 | -24.2375
14.8822 | -34.7709 | -24.2458
14.8883 | -34.7713 | -24.2537
14.8944 | -34.7718 | -24.2619
14.9005 | -34.7723 | -24.2699
14.9066 | -34.7727 | -24.2781
14.9127 | -34.7732 | -24.2861
14.9188 | -34.7737 | -24.2944
14.9249 | -34.7741 | -24.3025
14.931 | -34.7746 | -24.3106
14.9371 | -34.775 | -24.3186
14.9432 | -34.7755 | -24.3268
14.9493 | -34.776 | -24.3349
14.9554 | -34.7764 | -24.3429
14.9615 | -34.7768 | -24.3511
14.9676 | -34.7773 | -24.3594
14.9738 | -34.7778 | -24.3673
14.9799 | -34.7782 | -24.3758
14.986 | -34.7787 | -24.3835
14.9921 | -34.7792 | -24.3914
14.9982 | -34.7796 | -24.3999
15.0043 | -34.7801 | -24.4079
15.0104 | -34.7805 | -24.4161
15.0165 | -34.781 | -24.4243
15.0226 | -34.7815 | -24.4321
15.0287 | -34.7819 | -24.4407
15.0348 | -34.7824 | -24.4489
15.0409 | -34.7828 | -24.4572
15.047 | -34.7833 | -24.4656
15.0531 | -34.7839 | -24.4738
15.0592 | -34.7844 | -24.482
15.0653 | -34.7849 | -24.4901
15.0714 | -34.7853 | -24.4979
15.0775 | -34.7858 | -24.5057
15.0836 | -34.7863 | -24.5139
15.0897 | -34.7868 | -24.5218
15.0958 | -34.7872 | -24.5298
15.1019 | -34.7876 | -24.5378
15.108 | -34.7881 | -24.546
15.1141 | -34.7885 | -24.554
15.1203 | -34.7891 | -24.5621
15.1264 | -34.7896 | -24.5703
15.1325 | -34.7901 | -24.5783
15.1386 | -34.7905 | -24.5865
15.1447 | -34.791 | -24.5946
15.1508 | -34.7915 | -24.603
15.1569 | -34.792 | -24.6113
15.163 | -34.7926 | -24.6189
15.1691 | -34.793 | -24.6269
15.1752 | -34.7935 | -24.6349
15.1813 | -34.794 | -24.6431
15.1874 | -34.7945 | -24.6513
15.1935 | -34.795 | -24.6591
15.1996 | -34.7955 | -24.6674
15.2057 | -34.796 | -24.6753
15.2118 | -34.7964 | -24.6833
15.2179 | -34.7969 | -24.6911
15.224 | -34.7974 | -24.6993
15.2301 | -34.7979 | -24.7071
15.2362 | -34.7984 | -24.7151
15.2423 | -34.7988 | -24.7235
15.2484 | -34.7993 | -24.7314
15.2545 | -34.7998 | -24.7394
15.2607 | -34.8003 | -24.7475
15.2668 | -34.8008 | -24.7553
15.2729 | -34.8013 | -24.7632
15.279 | -34.8018 | -24.7714
15.2851 | -34.8023 | -24.7794
15.2912 | -34.8027 | -24.7873
15.2973 | -34.8032 | -24.7952
15.3034 | -34.8037 | -24.8031
15.3095 | -34.8041 | -24.8109
15.3156 | -34.8046 | -24.8191
15.3217 | -34.8051 | -24.8267
15.3278 | -34.8056 | -24.8348
15.3339 | -34.8061 | -24.8425
15.34 | -34.8066 | -24.8503
15.3461 | -34.807 | -24.8584
15.3522 | -34.8075 | -24.8663
15.3583 | -34.808 | -24.8742
15.3644 | -34.8085 | -24.8824
15.3705 | -34.809 | -24.8901
15.3766 | -34.8095 | -24.8983
15.3827 | -34.8099 | -24.906
15.3888 | -34.8104 | -24.9137
15.3949 | -34.8109 | -24.9215
15.401 | -34.8114 | -24.9295
15.4072 | -34.8119 | -24.9372
15.4133 | -34.8124 | -24.9452
15.4194 | -34.8129 | -24.9532
15.4255 | -34.8133 | -24.9612
15.4316 | -34.8139 | -24.9692
15.4377 | -34.8143 | -24.977
15.4438 | -34.8149 | -24.9845
15.4499 | -34.8153 | -24.9924
15.456 | -34.8158 | -25.0004
15.4621 | -34.8163 | -25.0083
15.4682 | -34.8168 | -25.0164
15.4743 | -34.8173 | -25.0241
15.4804 | -34.8178 | -25.0322
15.4865 | -34.8182 | -25.0401
15.4926 | -34.8187 | -25.0481
15.4987 | -34.8192 | -25.0561
15.5048 | -34.8197 | -25.064
15.5109 | -34.8202 | -25.0716
15.517 | -34.8207 | -25.0796
15.5231 | -34.8212 | -25.0875
15.5292 | -34.8217 | -25.0953
15.5353 | -34.8222 | -25.1032
15.5414 | -34.8227 | -25.1113
15.5476 | -34.8232 | -25.1193
15.5537 | -34.8237 | -25.1272
15.5598 | -34.8242 | -25.1349
15.5659 | -34.8247 | -25.143
15.572 | -34.8252 | -25.1505
15.5781 | -34.8257 | -25.1585
15.5842 | -34.8262 | -25.1661
15.5903 | -34.8267 | -25.1736
15.5964 | -34.8272 | -25.1813
15.6025 | -34.8277 | -25.1891
15.6086 | -34.8282 | -25.1971
15.6147 | -34.8286 | -25.205
15.6208 | -34.8292 | -25.2126
15.6269 | -34.8297 | -25.2204
15.633 | -34.8302 | -25.2283
15.6391 | -34.8307 | -25.2364
15.6452 | -34.8312 | -25.2441
15.6513 | -34.8317 | -25.2522
15.6574 | -34.8322 | -25.2597
15.6635 | -34.8326 | -25.2675
15.6696 | -34.8332 | -25.2753
15.6757 | -34.8337 | -25.2835
15.6818 | -34.8341 | -25.2917
15.688 | -34.8347 | -25.2996
15.6941 | -34.8352 | -25.3074
15.7002 | -34.8357 | -25.3153
15.7063 | -34.8361 | -25.3232
15.7124 | -34.8366 | -25.3312
15.7185 | -34.8371 | -25.3389
15.7246 | -34.8376 | -25.3466
15.7307 | -34.8381 | -25.3546
15.7368 | -34.8386 | -25.3627
15.7429 | -34.8391 | -25.3704
15.749 | -34.8396 | -25.3783
15.7551 | -34.8401 | -25.3861
15.7612 | -34.8406 | -25.3934
15.7673 | -34.841 | -25.4011
15.7734 | -34.8415 | -25.4087
15.7795 | -34.842 | -25.4169
15.7856 | -34.8425 | -25.4245
15.7917 | -34.8431 | -25.4321
15.7978 | -34.8436 | -25.44
15.8039 | -34.844 | -25.4478
15.81 | -34.8445 | -25.4555
15.8161 | -34.845 | -25.4635
15.8222 | -34.8455 | -25.4713
15.8283 | -34.846 | -25.4789
15.8345 | -34.8465 | -25.4868
15.8406 | -34.8471 | -25.4945
15.8467 | -34.8476 | -25.5025
15.8528 | -34.8481 | -25.5102
15.8589 | -34.8486 | -25.5179
15.865 | -34.8491 | -25.5256
15.8711 | -34.8496 | -25.5337
15.8772 | -34.8501 | -25.5415
15.8833 | -34.8506 | -25.5497
15.8894 | -34.8511 | -25.5575
15.8955 | -34.8516 | -25.5654
15.9016 | -34.8522 | -25.5731
15.9077 | -34.8526 | -25.5808
15.9138 | -34.8531 | -25.5885
15.9199 | -34.8536 | -25.5964
15.926 | -34.8542 | -25.6045
15.9321 | -34.8547 | -25.6123
15.9382 | -34.8552 | -25.62
15.9443 | -34.8558 | -25.6279
15.9504 | -34.8563 | -25.6357
15.9565 | -34.8568 | -25.6435
15.9626 | -34.8573 | -25.6513
15.9687 | -34.8579 | -25.6588
15.9749 | -34.8584 | -25.6668
15.981 | -34.8589 | -25.6746
15.9871 | -34.8594 | -25.6824
15.9932 | -34.8599 | -25.6903
15.9993 | -34.8604 | -25.698
16.0054 | -34.8609 | -25.7055
16.0115 | -34.8615 | -25.7132
16.0176 | -34.862 | -25.7212
16.0237 | -34.8625 | -25.7292
16.0298 | -34.863 | -25.7364
16.0359 | -34.8635 | -25.7441
16.042 | -34.8641 | -25.7518
16.0481 | -34.8646 | -25.7596
16.0542 | -34.8651 | -25.7676
16.0603 | -34.8656 | -25.7754
16.0664 | -34.8661 | -25.7833
16.0725 | -34.8667 | -25.7908
16.0786 | -34.8672 | -25.7988
16.0847 | -34.8677 | -25.8063
16.0908 | -34.8682 | -25.814
16.0969 | -34.8687 | -25.8216
16.103 | -34.8692 | -25.8296
16.1091 | -34.8697 | -25.8372
16.1152 | -34.8702 | -25.845
16.1214 | -34.8707 | -25.8527
16.1275 | -34.8712 | -25.8604
16.1336 | -34.8718 | -25.8682
16.1397 | -34.8723 | -25.876
16.1458 | -34.8728 | -25.8837
16.1519 | -34.8733 | -25.8915
16.158 | -34.8738 | -25.8993
16.1641 | -34.8744 | -25.9071
16.1702 | -34.8749 | -25.915
16.1763 | -34.8754 | -25.9228
16.1824 | -34.8759 | -25.9302
16.1885 | -34.8764 | -25.9379
16.1946 | -34.8769 | -25.9456
16.2007 | -34.8774 | -25.9537
16.2068 | -34.8779 | -25.9614
16.2129 | -34.8784 | -25.9692
16.219 | -34.8789 | -25.9769
16.2251 | -34.8794 | -25.9847
16.2312 | -34.8799 | -25.9926
16.2373 | -34.8804 | -26.0003
16.2434 | -34.8809 | -26.0079
16.2495 | -34.8814 | -26.0154
16.2556 | -34.882 | -26.0232
16.2618 | -34.8825 | -26.0306
16.2679 | -34.883 | -26.0384
16.274 | -34.8835 | -26.0464
16.2801 | -34.884 | -26.0538
16.2862 | -34.8845 | -26.0615
16.2923 | -34.8851 | -26.0694
16.2984 | -34.8855 | -26.0773
16.3045 | -34.8861 | -26.085
16.3106 | -34.8866 | -26.0924
16.3167 | -34.8871 | -26.1002
16.3228 | -34.8876 | -26.108
16.3289 | -34.8881 | -26.1158
16.335 | -34.8886 | -26.1237
16.3411 | -34.8891 | -26.1313
16.3472 | -34.8896 | -26.1392
16.3533 | -34.8901 | -26.1469
16.3594 | -34.8906 | -26.1545
16.3655 | -34.8911 | -26.1624
16.3716 | -34.8916 | -26.1701
16.3777 | -34.8921 | -26.1778
16.3838 | -34.8926 | -26.1854
16.3899 | -34.8931 | -26.1932
16.396 | -34.8936 | -26.201
16.4021 | -34.8942 | -26.2086
16.4083 | -34.8947 | -26.2163
16.4144 | -34.8953 | -26.224
16.4205 | -34.8958 | -26.2318
16.4266 | -34.8963 | -26.2395
16.4327 | -34.8968 | -26.247
16.4388 | -34.8973 | -26.2548
16.4449 | -34.8978 | -26.2624
16.451 | -34.8983 | -26.2703
16.4571 | -34.8988 | -26.2783
16.4632 | -34.8993 | -26.2858
16.4693 | -34.8998 | -26.2936
16.4754 | -34.9003 | -26.3015
16.4815 | -34.9008 | -26.3093
16.4876 | -34.9014 | -26.3169
16.4937 | -34.9019 | -26.3245
16.4998 | -34.9025 | -26.3323
16.5059 | -34.903 | -26.34
16.512 | -34.9035 | -26.348
16.5181 | -34.9041 | -26.3559
16.5242 | -34.9045 | -26.3637
16.5303 | -34.905 | -26.3713
16.5364 | -34.9055 | -26.3789
16.5425 | -34.906 | -26.3867
16.5487 | -34.9065 | -26.3943
16.5548 | -34.907 | -26.402
16.5609 | -34.9075 | -26.4099
16.567 | -34.908 | -26.4176
16.5731 | -34.9086 | -26.4254
16.5792 | -34.9091 | -26.4332
16.5853 | -34.9096 | -26.4408
16.5914 | -34.9101 | -26.4486
16.5975 | -34.9106 | -26.4564
16.6036 | -34.9111 | -26.464
16.6097 | -34.9116 | -26.4717
16.6158 | -34.9121 | -26.4796
16.6219 | -34.9127 | -26.4871
16.628 | -34.9131 | -26.4949
16.6341 | -34.9137 | -26.5025
16.6402 | -34.9142 | -26.5098
16.6463 | -34.9147 | -26.5174
16.6524 | -34.9152 | -26.525
16.6585 | -34.9157 | -26.5328
16.6646 | -34.9162 | -26.5405
16.6707 | -34.9167 | -26.5481
16.6768 | -34.9172 | -26.5558
16.6829 | -34.9178 | -26.5634
16.689 | -34.9183 | -26.5712
16.6952 | -34.9188 | -26.5789
16.7013 | -34.9193 | -26.5867
16.7074 | -34.9198 | -26.5943
16.7135 | -34.9203 | -26.6019
16.7196 | -34.9208 | -26.6097
16.7257 | -34.9213 | -26.6173
16.7318 | -34.9218 | -26.6249
16.7379 | -34.9223 | -26.6328
16.744 | -34.9228 | -26.6401
16.7501 | -34.9233 | -26.6479
16.7562 | -34.9238 | -26.6555
16.7623 | -34.9243 | -26.6631
16.7684 | -34.9249 | -26.671
16.7745 | -34.9254 | -26.6784
16.7806 | -34.9259 | -26.6862
16.7867 | -34.9264 | -26.6938
16.7928 | -34.9269 | -26.7015
16.7989 | -34.9274 | -26.709
16.805 | -34.9278 | -26.7168
16.8111 | -34.9284 | -26.7244
16.8172 | -34.9288 | -26.7324
16.8233 | -34.9294 | -26.7404
16.8294 | -34.9299 | -26.7482
16.8356 | -34.9304 | -26.7561
16.8417 | -34.9309 | -26.7636
16.8478 | -34.9314 | -26.7712
16.8539 | -34.9319 | -26.7789
16.86 | -34.9324 | -26.7865
16.8661 | -34.9329 | -26.7941
16.8722 | -34.9335 | -26.8019
16.8783 | -34.934 | -26.8094
16.8844 | -34.9345 | -26.8169
16.8905 | -34.935 | -26.8247
16.8966 | -34.9355 | -26.8325
16.9027 | -34.936 | -26.84
16.9088 | -34.9365 | -26.8481
16.9149 | -34.937 | -26.8561
16.921 | -34.9375 | -26.8636
16.9271 | -34.9381 | -26.8716
16.9332 | -34.9386 | -26.8792
16.9393 | -34.9391 | -26.8872
16.9454 | -34.9395 | -26.8947
16.9515 | -34.94 | -26.9026
16.9576 | -34.9405 | -26.9102
16.9637 | -34.9411 | -26.9179
16.9698 | -34.9416 | -26.9253
16.9759 | -34.9421 | -26.9328
16.9821 | -34.9426 | -26.9406
16.9882 | -34.9431 | -26.9484
16.9943 | -34.9436 | -26.9563
17.0004 | -34.9441 | -26.964
17.0065 | -34.9446 | -26.9717
17.0126 | -34.9451 | -26.9794
17.0187 | -34.9456 | -26.9872
17.0248 | -34.9461 | -26.9947
17.0309 | -34.9466 | -27.0023
17.037 | -34.9471 | -27.0101
17.0431 | -34.9476 | -27.0175
17.0492 | -34.9481 | -27.0252
17.0553 | -34.9485 | -27.0328
17.0614 | -34.949 | -27.0404
17.0675 | -34.9495 | -27.0484
17.0736 | -34.95 | -27.0562
17.0797 | -34.9505 | -27.0634
17.0858 | -34.951 | -27.0712
17.0919 | -34.9515 | -27.079
17.098 | -34.952 | -27.0865
17.1041 | -34.9525 | -27.0939
17.1102 | -34.953 | -27.1018
17.1163 | -34.9536 | -27.1095
17.1225 | -34.9541 | -27.117
17.1286 | -34.9546 | -27.1244
17.1347 | -34.9551 | -27.1321
17.1408 | -34.9556 | -27.1398
17.1469 | -34.9561 | -27.1473
17.153 | -34.9566 | -27.1549
17.1591 | -34.9571 | -27.1623
17.1652 | -34.9576 | -27.1698
17.1713 | -34.9581 | -27.1775
17.1774 | -34.9586 | -27.1853
17.1835 | -34.9591 | -27.1931
17.1896 | -34.9596 | -27.2006
17.1957 | -34.96 | -27.2083
17.2018 | -34.9605 | -27.2161
17.2079 | -34.961 | -27.2236
17.214 | -34.9615 | -27.2313
17.2201 | -34.962 | -27.2389
17.2262 | -34.9625 | -27.2466
17.2323 | -34.963 | -27.2542
17.2384 | -34.9635 | -27.2618
17.2445 | -34.9639 | -27.2693
17.2506 | -34.9644 | -27.2771
17.2567 | -34.9649 | -27.2848
17.2628 | -34.9654 | -27.2928
17.269 | -34.9659 | -27.3004
17.2751 | -34.9664 | -27.308
17.2812 | -34.9669 | -27.3155
17.2873 | -34.9674 | -27.3233
17.2934 | -34.9679 | -27.3309
17.2995 | -34.9683 | -27.3384
17.3056 | -34.9688 | -27.346
17.3117 | -34.9693 | -27.3537
17.3178 | -34.9698 | -27.3613
17.3239 | -34.9703 | -27.3688
17.33 | -34.9708 | -27.3763
17.3361 | -34.9712 | -27.3839
17.3422 | -34.9717 | -27.3916
17.3483 | -34.9722 | -27.3993
17.3544 | -34.9727 | -27.4068
17.3605 | -34.9732 | -27.4144
17.3666 | -34.9737 | -27.4221
17.3727 | -34.9742 | -27.4296
17.3788 | -34.9747 | -27.4371
17.3849 | -34.9752 | -27.4446
17.391 | -34.9757 | -27.452
17.3971 | -34.9762 | -27.4597
17.4032 | -34.9767 | -27.4676
17.4094 | -34.9771 | -27.4752
17.4155 | -34.9776 | -27.4826
17.4216 | -34.9781 | -27.4903
17.4277 | -34.9786 | -27.4981
17.4338 | -34.9791 | -27.5057
17.4399 | -34.9796 | -27.5134
17.446 | -34.9801 | -27.5209
17.4521 | -34.9806 | -27.5284
17.4582 | -34.9811 | -27.5359
17.4643 | -34.9815 | -27.5435
17.4704 | -34.982 | -27.5509
17.4765 | -34.9825 | -27.5586
17.4826 | -34.983 | -27.5662
17.4887 | -34.9835 | -27.5741
17.4948 | -34.9839 | -27.5816
17.5009 | -34.9844 | -27.589
17.507 | -34.9849 | -27.5967
17.5131 | -34.9854 | -27.6043
17.5192 | -34.9858 | -27.6119
17.5253 | -34.9863 | -27.6195
17.5314 | -34.9867 | -27.6271
17.5375 | -34.9872 | -27.6348
17.5436 | -34.9877 | -27.6424
17.5497 | -34.9882 | -27.6502
17.5559 | -34.9886 | -27.6576
17.562 | -34.9891 | -27.6651
17.5681 | -34.9896 | -27.6726
17.5742 | -34.9901 | -27.6802
17.5803 | -34.9906 | -27.6876
17.5864 | -34.9911 | -27.695
17.5925 | -34.9916 | -27.7028
17.5986 | -34.9921 | -27.7101
17.6047 | -34.9926 | -27.7181
17.6108 | -34.9931 | -27.7254
17.6169 | -34.9936 | -27.7329
17.623 | -34.994 | -27.7404
17.6291 | -34.9946 | -27.7479
17.6352 | -34.995 | -27.7553
17.6413 | -34.9955 | -27.7629
17.6474 | -34.996 | -27.7706
17.6535 | -34.9965 | -27.7783
17.6596 | -34.9969 | -27.7858
17.6657 | -34.9974 | -27.7933
17.6718 | -34.9979 | -27.801
17.6779 | -34.9984 | -27.8085
17.684 | -34.9989 | -27.816
17.6901 | -34.9994 | -27.8235
17.6963 | -34.9998 | -27.8309
17.7024 | -35.0003 | -27.8384
17.7085 | -35.0008 | -27.8457
17.7146 | -35.0013 | -27.8532
17.7207 | -35.0017 | -27.8607
17.7268 | -35.0022 | -27.8682
17.7329 | -35.0027 | -27.8759
17.739 | -35.0032 | -27.8836
17.7451 | -35.0037 | -27.891
17.7512 | -35.0042 | -27.8986
17.7573 | -35.0046 | -27.9062
17.7634 | -35.0051 | -27.9135
17.7695 | -35.0056 | -27.9211
17.7756 | -35.0061 | -27.9286
17.7817 | -35.0066 | -27.9359
17.7878 | -35.0071 | -27.9434
17.7939 | -35.0075 | -27.9509
17.8 | -35.008 | -27.9584
17.8061 | -35.0085 | -27.966
17.8122 | -35.009 | -27.9735
17.8183 | -35.0095 | -27.9812
17.8244 | -35.01 | -27.9887
17.8305 | -35.0104 | -27.9962
17.8366 | -35.0109 | -28.0034
17.8428 | -35.0114 | -28.0111
17.8489 | -35.0119 | -28.0185
17.855 | -35.0124 | -28.0258
17.8611 | -35.0128 | -28.0332
17.8672 | -35.0133 | -28.0404
17.8733 | -35.0138 | -28.0482
17.8794 | -35.0143 | -28.0558
17.8855 | -35.0148 | -28.0632
17.8916 | -35.0153 | -28.0706
17.8977 | -35.0158 | -28.0781
17.9038 | -35.0162 | -28.0854
17.9099 | -35.0167 | -28.0929
17.916 | -35.0172 | -28.1004
17.9221 | -35.0176 | -28.1079
17.9282 | -35.0181 | -28.1154
17.9343 | -35.0186 | -28.1231
17.9404 | -35.0191 | -28.1304
17.9465 | -35.0196 | -28.1381
17.9526 | -35.02 | -28.1458
17.9587 | -35.0205 | -28.1532
17.9648 | -35.021 | -28.1607
17.9709 | -35.0215 | -28.1681
17.977 | -35.0219 | -28.1755
17.9832 | -35.0224 | -28.1831
17.9893 | -35.0229 | -28.1906
17.9954 | -35.0233 | -28.1984
18.0015 | -35.0239 | -28.206
18.0076 | -35.0243 | -28.213
18.0137 | -35.0248 | -28.2207
18.0198 | -35.0252 | -28.2282
18.0259 | -35.0257 | -28.2358
18.032 | -35.0261 | -28.243
18.0381 | -35.0266 | -28.2503
18.0442 | -35.0271 | -28.2579
18.0503 | -35.0276 | -28.2655
18.0564 | -35.028 | -28.2729
18.0625 | -35.0285 | -28.2805
18.0686 | -35.0289 | -28.2878
18.0747 | -35.0294 | -28.2952
18.0808 | -35.0299 | -28.3026
18.0869 | -35.0304 | -28.3101
18.093 | -35.0309 | -28.3176
18.0991 | -35.0313 | -28.3254
18.1052 | -35.0318 | -28.3331
18.1113 | -35.0323 | -28.3403
18.1174 | -35.0328 | -28.3476
18.1236 | -35.0333 | -28.3552
18.1297 | -35.0338 | -28.3626
18.1358 | -35.0343 | -28.3701
18.1419 | -35.0348 | -28.3776
18.148 | -35.0353 | -28.385
18.1541 | -35.0358 | -28.3923
18.1602 | -35.0363 | -28.3998
18.1663 | -35.0368 | -28.4072
18.1724 | -35.0373 | -28.4149
18.1785 | -35.0378 | -28.4225
18.1846 | -35.0382 | -28.4297
18.1907 | -35.0387 | -28.4375
18.1968 | -35.0391 | -28.4452
18.2029 | -35.0396 | -28.4523
18.209 | -35.0401 | -28.46
18.2151 | -35.0406 | -28.4678
18.2212 | -35.0411 | -28.475
18.2273 | -35.0417 | -28.4824
18.2334 | -35.0421 | -28.4898
18.2395 | -35.0426 | -28.4971
18.2456 | -35.0431 | -28.5044
18.2517 | -35.0435 | -28.5118
18.2578 | -35.044 | -28.5191
18.2639 | -35.0445 | -28.5269
18.2701 | -35.045 | -28.5343
18.2762 | -35.0455 | -28.5418
18.2823 | -35.046 | -28.5494
18.2884 | -35.0465 | -28.5571
18.2945 | -35.047 | -28.5643
18.3006 | -35.0475 | -28.5718
18.3067 | -35.048 | -28.5792
18.3128 | -35.0484 | -28.5866
18.3189 | -35.0489 | -28.5942
18.325 | -35.0494 | -28.6012
18.3311 | -35.0499 | -28.6088
18.3372 | -35.0503 | -28.616
18.3433 | -35.0508 | -28.6232
18.3494 | -35.0512 | -28.6306
18.3555 | -35.0517 | -28.6378
18.3616 | -35.0522 | -28.645
18.3677 | -35.0527 | -28.6523
18.3738 | -35.0532 | -28.6596
18.3799 | -35.0536 | -28.6668
18.386 | -35.0541 | -28.674
18.3921 | -35.0546 | -28.6813
18.3982 | -35.0551 | -28.6885
18.4043 | -35.0555 | -28.6959
18.4105 | -35.056 | -28.7033
18.4166 | -35.0565 | -28.7106
18.4227 | -35.0569 | -28.718
18.4288 | -35.0574 | -28.7255
18.4349 | -35.0579 | -28.7328
18.441 | -35.0584 | -28.7402
18.4471 | -35.0589 | -28.7475
18.4532 | -35.0594 | -28.755
18.4593 | -35.0598 | -28.7623
18.4654 | -35.0603 | -28.7694
18.4715 | -35.0608 | -28.7767
18.4776 | -35.0613 | -28.7841
18.4837 | -35.0618 | -28.7916
18.4898 | -35.0623 | -28.7992
18.4959 | -35.0628 | -28.8066
18.502 | -35.0633 | -28.8141
18.5081 | -35.0637 | -28.8216
18.5142 | -35.0641 | -28.8292
18.5203 | -35.0646 | -28.8363
18.5264 | -35.0651 | -28.8434
18.5325 | -35.0656 | -28.851
18.5386 | -35.0661 | -28.8586
18.5447 | -35.0666 | -28.8659
18.5508 | -35.0671 | -28.873
18.557 | -35.0676 | -28.8805
18.5631 | -35.068 | -28.8879
18.5692 | -35.0685 | -28.8952
18.5753 | -35.069 | -28.9024
18.5814 | -35.0695 | -28.91
18.5875 | -35.07 | -28.9176
18.5936 | -35.0704 | -28.9252
18.5997 | -35.0709 | -28.9327
18.6058 | -35.0714 | -28.94
18.6119 | -35.0718 | -28.9475
18.618 | -35.0724 | -28.9548
18.6241 | -35.0728 | -28.9624
18.6302 | -35.0734 | -28.9701
18.6363 | -35.0738 | -28.9772
18.6424 | -35.0743 | -28.9844
18.6485 | -35.0748 | -28.9918
18.6546 | -35.0753 | -28.999
18.6607 | -35.0757 | -29.0063
18.6668 | -35.0763 | -29.0137
18.6729 | -35.0767 | -29.0213
18.679 | -35.0772 | -29.0282
18.6851 | -35.0777 | -29.0355
18.6912 | -35.0782 | -29.0428
18.6974 | -35.0786 | -29.0501
18.7035 | -35.0791 | -29.0574
18.7096 | -35.0796 | -29.0648
18.7157 | -35.0801 | -29.072
18.7218 | -35.0806 | -29.0793
18.7279 | -35.0811 | -29.0865
18.734 | -35.0816 | -29.0938
18.7401 | -35.0821 | -29.1009
18.7462 | -35.0826 | -29.1079
18.7523 | -35.0831 | -29.115
18.7584 | -35.0836 | -29.1226
18.7645 | -35.0841 | -29.1299
18.7706 | -35.0846 | -29.1372
18.7767 | -35.0851 | -29.1443
18.7828 | -35.0856 | -29.1511
18.7889 | -35.0861 | -29.1583
18.795 | -35.0866 | -29.1654
18.8011 | -35.0871 | -29.1727
18.8072 | -35.0876 | -29.1802
18.8133 | -35.088 | -29.1877
18.8194 | -35.0885 | -29.1946
18.8255 | -35.089 | -29.2017
18.8316 | -35.0894 | -29.209
18.8377 | -35.0899 | -29.2162
18.8439 | -35.0904 | -29.2234
18.85 | -35.0909 | -29.2308
18.8561 | -35.0914 | -29.238
18.8622 | -35.0918 | -29.2455
18.8683 | -35.0923 | -29.2525
18.8744 | -35.0928 | -29.2598
18.8805 | -35.0933 | -29.2671
18.8866 | -35.0938 | -29.2746
18.8927 | -35.0943 | -29.2816
18.8988 | -35.0948 | -29.2887
18.9049 | -35.0953 | -29.2959
18.911 | -35.0958 | -29.3032
18.9171 | -35.0963 | -29.3105
18.9232 | -35.0968 | -29.3175
18.9293 | -35.0973 | -29.3248
18.9354 | -35.0977 | -29.3319
18.9415 | -35.0982 | -29.3389
18.9476 | -35.0987 | -29.3462
18.9537 | -35.0992 | -29.3537
18.9598 | -35.0997 | -29.3606
18.9659 | -35.1002 | -29.3678
18.972 | -35.1007 | -29.3751
18.9781 | -35.1012 | -29.3822
18.9843 | -35.1017 | -29.3895
18.9904 | -35.1021 | -29.3969
18.9965 | -35.1026 | -29.4041
19.0026 | -35.1031 | -29.4114
19.0087 | -35.1036 | -29.4189
19.0148 | -35.1041 | -29.426
19.0209 | -35.1046 | -29.4336
19.027 | -35.1051 | -29.441
19.0331 | -35.1056 | -29.4482
19.0392 | -35.1061 | -29.4554
19.0453 | -35.1066 | -29.4626
19.0514 | -35.1071 | -29.4698
19.0575 | -35.1075 | -29.4771
19.0636 | -35.108 | -29.4838
19.0697 | -35.1085 | -29.4912
19.0758 | -35.109 | -29.4984
19.0819 | -35.1095 | -29.5058
19.088 | -35.11 | -29.5127
19.0941 | -35.1105 | -29.5196
19.1002 | -35.1109 | -29.5265
19.1063 | -35.1114 | -29.5338
19.1124 | -35.112 | -29.541
19.1185 | -35.1124 | -29.5484
19.1246 | -35.1129 | -29.5556
19.1308 | -35.1134 | -29.5624
19.1369 | -35.1139 | -29.5693
19.143 | -35.1144 | -29.5765
19.1491 | -35.1149 | -29.5836
19.1552 | -35.1153 | -29.5912
19.1613 | -35.1158 | -29.5983
19.1674 | -35.1163 | -29.6057
19.1735 | -35.1168 | -29.6129
19.1796 | -35.1173 | -29.6198
19.1857 | -35.1178 | -29.6271
19.1918 | -35.1183 | -29.6345
19.1979 | -35.1187 | -29.6418
19.204 | -35.1193 | -29.6488
19.2101 | -35.1198 | -29.6561
19.2162 | -35.1202 | -29.6636
19.2223 | -35.1207 | -29.6704
19.2284 | -35.1212 | -29.6778
19.2345 | -35.1217 | -29.6851
19.2406 | -35.1222 | -29.6923
19.2467 | -35.1227 | -29.6995
19.2528 | -35.1232 | -29.7067
19.2589 | -35.1236 | -29.7139
19.265 | -35.1241 | -29.7211
19.2712 | -35.1246 | -29.7284
19.2773 | -35.1251 | -29.7354
19.2834 | -35.1256 | -29.7422
19.2895 | -35.1261 | -29.7496
19.2956 | -35.1266 | -29.757
19.3017 | -35.1271 | -29.7641
19.3078 | -35.1276 | -29.7712
19.3139 | -35.1281 | -29.7782
19.32 | -35.1286 | -29.7852
19.3261 | -35.1291 | -29.7923
19.3322 | -35.1296 | -29.7995
19.3383 | -35.1301 | -29.8063
19.3444 | -35.1306 | -29.8132
19.3505 | -35.1311 | -29.8203
19.3566 | -35.1316 | -29.8274
19.3627 | -35.132 | -29.8347
19.3688 | -35.1325 | -29.8421
19.3749 | -35.133 | -29.8494
19.381 | -35.1335 | -29.8563
19.3871 | -35.134 | -29.8634
19.3932 | -35.1345 | -29.8709
19.3993 | -35.1349 | -29.8779
19.4054 | -35.1354 | -29.8852
19.4115 | -35.1359 | -29.8923
19.4177 | -35.1364 | -29.8994
19.4238 | -35.1369 | -29.9064
19.4299 | -35.1374 | -29.9133
19.436 | -35.1379 | -29.9208
19.4421 | -35.1384 | -29.9275
19.4482 | -35.1389 | -29.9347
19.4543 | -35.1394 | -29.9415
19.4604 | -35.1399 | -29.9486
19.4665 | -35.1404 | -29.9556
19.4726 | -35.1409 | -29.9624
19.4787 | -35.1414 | -29.9695
19.4848 | -35.1419 | -29.9765
19.4909 | -35.1424 | -29.9836
19.497 | -35.1428 | -29.9907
19.5031 | -35.1433 | -29.9976
19.5092 | -35.1438 | -30.0048
19.5153 | -35.1443 | -30.0119
19.5214 | -35.1448 | -30.0187
19.5275 | -35.1453 | -30.0256
19.5336 | -35.1457 | -30.0327
19.5397 | -35.1462 | -30.0398
19.5458 | -35.1468 | -30.0469
19.5519 | -35.1472 | -30.0539
19.5581 | -35.1477 | -30.0608
19.5642 | -35.1482 | -30.0677
19.5703 | -35.1487 | -30.0747
19.5764 | -35.1492 | -30.0819
19.5825 | -35.1497 | -30.089
19.5886 | -35.1501 | -30.096
19.5947 | -35.1506 | -30.1028
19.6008 | -35.1511 | -30.1098
19.6069 | -35.1516 | -30.1168
19.613 | -35.1521 | -30.1239
19.6191 | -35.1526 | -30.131
19.6252 | -35.1531 | -30.1379
19.6313 | -35.1535 | -30.1448
19.6374 | -35.154 | -30.1518
19.6435 | -35.1545 | -30.1587
19.6496 | -35.155 | -30.1657
19.6557 | -35.1555 | -30.1725
19.6618 | -35.156 | -30.1793
19.6679 | -35.1565 | -30.1862
19.674 | -35.1569 | -30.1932
19.6801 | -35.1574 | -30.2002
19.6862 | -35.1579 | -30.207
19.6923 | -35.1584 | -30.2138
19.6984 | -35.1589 | -30.2205
19.7046 | -35.1594 | -30.2276
19.7107 | -35.1599 | -30.2346
19.7168 | -35.1604 | -30.2418
19.7229 | -35.1608 | -30.249
19.729 | -35.1613 | -30.2561
19.7351 | -35.1618 | -30.2626
19.7412 | -35.1623 | -30.2696
19.7473 | -35.1628 | -30.2767
19.7534 | -35.1633 | -30.2838
19.7595 | -35.1638 | -30.291
19.7656 | -35.1643 | -30.2984
19.7717 | -35.1648 | -30.3049
19.7778 | -35.1652 | -30.3122
19.7839 | -35.1657 | -30.3189
19.79 | -35.1662 | -30.3261
19.7961 | -35.1667 | -30.3332
19.8022 | -35.1672 | -30.3404
19.8083 | -35.1677 | -30.3474
19.8144 | -35.1682 | -30.3543
19.8205 | -35.1686 | -30.3614
19.8266 | -35.1691 | -30.3684
19.8327 | -35.1696 | -30.3753
19.8388 | -35.1701 | -30.3826
19.845 | -35.1706 | -30.3894
19.8511 | -35.1711 | -30.3968
19.8572 | -35.1716 | -30.4038
19.8633 | -35.172 | -30.4107
19.8694 | -35.1726 | -30.418
19.8755 | -35.1731 | -30.4248
19.8816 | -35.1736 | -30.4318
19.8877 | -35.1741 | -30.4391
19.8938 | -35.1746 | -30.4458
19.8999 | -35.1751 | -30.4528
19.906 | -35.1756 | -30.4598
19.9121 | -35.1762 | -30.467
19.9182 | -35.1767 | -30.4736
19.9243 | -35.1772 | -30.4808
19.9304 | -35.1776 | -30.4881
19.9365 | -35.1781 | -30.4951
19.9426 | -35.1785 | -30.5018
19.9487 | -35.1791 | -30.5085
19.9548 | -35.1796 | -30.5155
19.9609 | -35.1801 | -30.5225
19.967 | -35.1806 | -30.5293
19.9731 | -35.1811 | -30.5362
19.9792 | -35.1816 | -30.5428
19.9853 | -35.1821 | -30.5496
19.9915 | -35.1826 | -30.5568
19.9976 | -35.1831 | -30.5635
20.0037 | -35.1836 | -30.5703
20.0098 | -35.1841 | -30.5773
20.0159 | -35.1846 | -30.5844
20.022 | -35.1851 | -30.5917
20.0281 | -35.1856 | -30.5989
20.0342 | -35.1861 | -30.6057
20.0403 | -35.1866 | -30.6124
20.0464 | -35.187 | -30.6201
20.0525 | -35.1875 | -30.6267
20.0586 | -35.188 | -30.6335
20.0647 | -35.1885 | -30.6405
20.0708 | -35.189 | -30.6479
20.0769 | -35.1895 | -30.6548
20.083 | -35.19 | -30.6615
20.0891 | -35.1904 | -30.6685
20.0952 | -35.191 | -30.6753
20.1013 | -35.1914 | -30.6822
20.1074 | -35.1919 | -30.6891
20.1135 | -35.1924 | -30.6958
20.1196 | -35.1929 | -30.7027
20.1257 | -35.1934 | -30.7098
20.1319 | -35.1939 | -30.7168
20.138 | -35.1944 | -30.7234
20.1441 | -35.1948 | -30.7303
20.1502 | -35.1953 | -30.7372
20.1563 | -35.1959 | -30.7436
20.1624 | -35.1963 | -30.7508
20.1685 | -35.1968 | -30.7579
20.1746 | -35.1973 | -30.7652
20.1807 | -35.1978 | -30.7723
20.1868 | -35.1982 | -30.7793
20.1929 | -35.1988 | -30.7855
20.199 | -35.1992 | -30.7924
20.2051 | -35.1997 | -30.799
20.2112 | -35.2001 | -30.8058
20.2173 | -35.2005 | -30.8125
20.2234 | -35.201 | -30.8197
20.2295 | -35.2015 | -30.8266
20.2356 | -35.202 | -30.834
20.2417 | -35.2024 | -30.8409
20.2478 | -35.2028 | -30.8479
20.2539 | -35.2033 | -30.8551
20.26 | -35.2038 | -30.8618
20.2661 | -35.2043 | -30.869
20.2723 | -35.2047 | -30.8758
20.2784 | -35.2052 | -30.8825
20.2845 | -35.2057 | -30.8895
20.2906 | -35.2062 | -30.8966
20.2967 | -35.2067 | -30.9033
20.3028 | -35.2072 | -30.9098
20.3089 | -35.2076 | -30.917
20.315 | -35.2081 | -30.9236
20.3211 | -35.2085 | -30.9304
20.3272 | -35.2091 | -30.9371
20.3333 | -35.2095 | -30.9436
20.3394 | -35.2099 | -30.9505
20.3455 | -35.2105 | -30.9575
20.3516 | -35.211 | -30.9644
20.3577 | -35.2114 | -30.9712
20.3638 | -35.2119 | -30.9781
20.3699 | -35.2124 | -30.9846
20.376 | -35.213 | -30.9916
20.3821 | -35.2135 | -30.9985
20.3882 | -35.2139 | -31.005
20.3943 | -35.2144 | -31.012
20.4004 | -35.2149 | -31.0184
20.4065 | -35.2154 | -31.025
20.4126 | -35.2159 | -31.0318
20.4188 | -35.2164 | -31.0389
20.4249 | -35.2169 | -31.0459
20.431 | -35.2173 | -31.0527
20.4371 | -35.2178 | -31.059
20.4432 | -35.2183 | -31.0659
20.4493 | -35.2188 | -31.0727
20.4554 | -35.2193 | -31.0795
20.4615 | -35.2198 | -31.0867
20.4676 | -35.2203 | -31.0936
20.4737 | -35.2207 | -31.1001
20.4798 | -35.2213 | -31.107
20.4859 | -35.2218 | -31.1137
20.492 | -35.2223 | -31.1204
20.4981 | -35.2227 | -31.1274
20.5042 | -35.2232 | -31.134
20.5103 | -35.2237 | -31.1407
20.5164 | -35.2242 | -31.1474
20.5225 | -35.2246 | -31.1542
20.5286 | -35.2251 | -31.1613
20.5347 | -35.2256 | -31.168
20.5408 | -35.2261 | -31.1749
20.5469 | -35.2266 | -31.1814
20.553 | -35.2271 | -31.1884
20.5592 | -35.2276 | -31.1952
20.5653 | -35.228 | -31.2022
20.5714 | -35.2285 | -31.2085
20.5775 | -35.229 | -31.2152
20.5836 | -35.2294 | -31.2222
20.5897 | -35.23 | -31.2289
20.5958 | -35.2304 | -31.2357
20.6019 | -35.2309 | -31.2423
20.608 | -35.2314 | -31.2493
20.6141 | -35.2318 | -31.256
20.6202 | -35.2323 | -31.2628
20.6263 | -35.2327 | -31.2692
20.6324 | -35.2332 | -31.2759
20.6385 | -35.2337 | -31.2826
20.6446 | -35.2341 | -31.2896
20.6507 | -35.2346 | -31.2962
20.6568 | -35.2351 | -31.3029
20.6629 | -35.2356 | -31.3094
20.669 | -35.236 | -31.316
20.6751 | -35.2365 | -31.3227
20.6812 | -35.237 | -31.3292
20.6873 | -35.2375 | -31.3361
20.6934 | -35.238 | -31.3429
20.6995 | -35.2385 | -31.3495
20.7057 | -35.2389 | -31.3562
20.7118 | -35.2394 | -31.3629
20.7179 | -35.2399 | -31.3695
20.724 | -35.2404 | -31.3761
20.7301 | -35.2408 | -31.3831
20.7362 | -35.2413 | -31.3898
20.7423 | -35.2418 | -31.3964
20.7484 | -35.2423 | -31.4032
20.7545 | -35.2428 | -31.4097
20.7606 | -35.2433 | -31.4168
20.7667 | -35.2437 | -31.4236
20.7728 | -35.2442 | -31.4304
20.7789 | -35.2447 | -31.4373
20.785 | -35.2452 | -31.4443
20.7911 | -35.2457 | -31.4507
20.7972 | -35.2461 | -31.4575
20.8033 | -35.2465 | -31.4644
20.8094 | -35.2471 | -31.4716
20.8155 | -35.2476 | -31.4786
20.8216 | -35.248 | -31.4854
20.8277 | -35.2485 | -31.4917
20.8338 | -35.249 | -31.4982
20.8399 | -35.2495 | -31.5051
20.8461 | -35.25 | -31.5114
20.8522 | -35.2505 | -31.5182
20.8583 | -35.251 | -31.5246
20.8644 | -35.2514 | -31.5313
20.8705 | -35.2519 | -31.5382
20.8766 | -35.2524 | -31.5452
20.8827 | -35.2529 | -31.552
20.8888 | -35.2533 | -31.5585
20.8949 | -35.2538 | -31.5652
20.901 | -35.2543 | -31.5718
20.9071 | -35.2547 | -31.5785
20.9132 | -35.2552 | -31.5852
20.9193 | -35.2557 | -31.592
20.9254 | -35.2561 | -31.5985
20.9315 | -35.2566 | -31.6045
20.9376 | -35.257 | -31.6113
20.9437 | -35.2575 | -31.6182
20.9498 | -35.2579 | -31.6248
20.9559 | -35.2584 | -31.6314
20.962 | -35.2589 | -31.6382
20.9681 | -35.2593 | -31.6451
20.9742 | -35.2598 | -31.6518
20.9803 | -35.2602 | -31.6585
20.9864 | -35.2607 | -31.665
20.9926 | -35.2612 | -31.6721
20.9987 | -35.2616 | -31.6783
21.0048 | -35.2621 | -31.6851
21.0109 | -35.2626 | -31.6916
21.017 | -35.263 | -31.6984
21.0231 | -35.2635 | -31.705
21.0292 | -35.264 | -31.7116
21.0353 | -35.2644 | -31.7183
21.0414 | -35.2649 | -31.7249
21.0475 | -35.2654 | -31.7318
21.0536 | -35.2658 | -31.7384
21.0597 | -35.2663 | -31.7448
21.0658 | -35.2667 | -31.7513
21.0719 | -35.2673 | -31.7583
21.078 | -35.2677 | -31.7652
21.0841 | -35.2681 | -31.7713
21.0902 | -35.2686 | -31.7779
21.0963 | -35.2691 | -31.7847
21.1024 | -35.2696 | -31.7916
21.1085 | -35.27 | -31.798
21.1146 | -35.2704 | -31.8049
21.1207 | -35.2709 | -31.8114
21.1268 | -35.2714 | -31.8178
21.133 | -35.2719 | -31.8246
21.1391 | -35.2723 | -31.8307
21.1452 | -35.2728 | -31.8374
21.1513 | -35.2732 | -31.844
21.1574 | -35.2737 | -31.8509
21.1635 | -35.2741 | -31.8572
21.1696 | -35.2746 | -31.8636
21.1757 | -35.275 | -31.8698
21.1818 | -35.2755 | -31.8765
21.1879 | -35.2759 | -31.8831
21.194 | -35.2764 | -31.8895
21.2001 | -35.2769 | -31.896
21.2062 | -35.2773 | -31.9028
21.2123 | -35.2778 | -31.9093
21.2184 | -35.2782 | -31.9158
21.2245 | -35.2787 | -31.9224
21.2306 | -35.2791 | -31.9284
21.2367 | -35.2796 | -31.935
21.2428 | -35.2801 | -31.9416
21.2489 | -35.2805 | -31.9477
21.255 | -35.281 | -31.9543
21.2611 | -35.2814 | -31.9609
21.2672 | -35.2819 | -31.9673
21.2733 | -35.2823 | -31.9738
21.2795 | -35.2828 | -31.9799
21.2856 | -35.2832 | -31.9865
21.2917 | -35.2837 | -31.9932
21.2978 | -35.2842 | -31.9998
21.3039 | -35.2846 | -32.0062
21.31 | -35.2851 | -32.0129
21.3161 | -35.2855 | -32.0195
21.3222 | -35.286 | -32.0256
21.3283 | -35.2864 | -32.0324
21.3344 | -35.2869 | -32.0388
21.3405 | -35.2873 | -32.0457
21.3466 | -35.2878 | -32.0523
21.3527 | -35.2882 | -32.0587
21.3588 | -35.2887 | -32.0649
21.3649 | -35.2891 | -32.0712
21.371 | -35.2896 | -32.0777
21.3771 | -35.29 | -32.0841
21.3832 | -35.2905 | -32.0903
21.3893 | -35.2909 | -32.0966
21.3954 | -35.2914 | -32.1032
21.4015 | -35.2918 | -32.1098
21.4076 | -35.2923 | -32.1164
21.4137 | -35.2928 | -32.1232
21.4199 | -35.2932 | -32.1297
21.426 | -35.2936 | -32.1363
21.4321 | -35.294 | -32.1427
21.4382 | -35.2945 | -32.1493
21.4443 | -35.2949 | -32.1557
21.4504 | -35.2953 | -32.1625
21.4565 | -35.2958 | -32.1691
21.4626 | -35.2963 | -32.1755
21.4687 | -35.2967 | -32.1819
21.4748 | -35.2972 | -32.1885
21.4809 | -35.2976 | -32.1948
21.487 | -35.2981 | -32.2013
21.4931 | -35.2986 | -32.2075
21.4992 | -35.299 | -32.2137
21.5053 | -35.2995 | -32.2204
21.5114 | -35.2999 | -32.2267
21.5175 | -35.3004 | -32.2333
21.5236 | -35.3008 | -32.2398
21.5297 | -35.3012 | -32.2459
21.5358 | -35.3016 | -32.2521
21.5419 | -35.3021 | -32.2588
21.548 | -35.3025 | -32.2652
21.5541 | -35.3029 | -32.2718
21.5602 | -35.3034 | -32.2784
21.5664 | -35.3038 | -32.2847
21.5725 | -35.3042 | -32.291
21.5786 | -35.3046 | -32.2974
21.5847 | -35.3051 | -32.3039
21.5908 | -35.3055 | -32.3103
21.5969 | -35.3059 | -32.3165
21.603 | -35.3064 | -32.3233
21.6091 | -35.3069 | -32.3302
21.6152 | -35.3073 | -32.3371
21.6213 | -35.3078 | -32.3438
21.6274 | -35.3083 | -32.3502
21.6335 | -35.3087 | -32.3564
21.6396 | -35.3091 | -32.3629
21.6457 | -35.3096 | -32.3696
21.6518 | -35.31 | -32.3761
21.6579 | -35.3105 | -32.3824
21.664 | -35.3109 | -32.389
21.6701 | -35.3114 | -32.3953
21.6762 | -35.3118 | -32.4018
21.6823 | -35.3123 | -32.4085
21.6884 | -35.3127 | -32.415
21.6945 | -35.3131 | -32.4212
21.7006 | -35.3135 | -32.4275
21.7068 | -35.314 | -32.4338
21.7129 | -35.3144 | -32.4403
21.719 | -35.3149 | -32.4468
21.7251 | -35.3153 | -32.4529
21.7312 | -35.3157 | -32.4593
21.7373 | -35.3162 | -32.4658
21.7434 | -35.3166 | -32.4725
21.7495 | -35.317 | -32.4786
21.7556 | -35.3175 | -32.4853
21.7617 | -35.318 | -32.4916
21.7678 | -35.3184 | -32.4981
21.7739 | -35.3189 | -32.5043
21.78 | -35.3193 | -32.5107
21.7861 | -35.3198 | -32.5172
21.7922 | -35.3202 | -32.5237
21.7983 | -35.3206 | -32.5302
21.8044 | -35.3211 | -32.5365
21.8105 | -35.3216 | -32.5427
21.8166 | -35.322 | -32.5491
21.8227 | -35.3224 | -32.5555
21.8288 | -35.3229 | -32.5616
21.8349 | -35.3233 | -32.5679
21.841 | -35.3238 | -32.5745
21.8471 | -35.3242 | -32.581
21.8533 | -35.3246 | -32.5873
21.8594 | -35.3251 | -32.594
21.8655 | -35.3256 | -32.6006
21.8716 | -35.326 | -32.6072
21.8777 | -35.3265 | -32.6137
21.8838 | -35.327 | -32.6201
21.8899 | -35.3274 | -32.6264
21.896 | -35.3279 | -32.6327
21.9021 | -35.3284 | -32.6393
21.9082 | -35.3288 | -32.6457
21.9143 | -35.3292 | -32.6518
21.9204 | -35.3296 | -32.658
21.9265 | -35.33 | -32.6643
21.9326 | -35.3304 | -32.6708
21.9387 | -35.3309 | -32.6773
21.9448 | -35.3314 | -32.6838
21.9509 | -35.3318 | -32.6898
21.957 | -35.3323 | -32.696
21.9631 | -35.3327 | -32.7028
21.9692 | -35.3331 | -32.7093
21.9753 | -35.3335 | -32.7156
21.9814 | -35.334 | -32.7222
21.9875 | -35.3344 | -32.7284
21.9937 | -35.3348 | -32.7349
21.9998 | -35.3353 | -32.7417
22.0059 | -35.3357 | -32.7481
22.012 | -35.3361 | -32.7543
22.0181 | -35.3366 | -32.7604
22.0242 | -35.337 | -32.7667
22.0303 | -35.3374 | -32.7732
22.0364 | -35.3378 | -32.7796
22.0425 | -35.3383 | -32.7862
22.0486 | -35.3387 | -32.7924
22.0547 | -35.3391 | -32.7984
22.0608 | -35.3395 | -32.8048
22.0669 | -35.34 | -32.811
22.073 | -35.3404 | -32.8175
22.0791 | -35.3409 | -32.8238
22.0852 | -35.3413 | -32.8302
22.0913 | -35.3418 | -32.8365
22.0974 | -35.3422 | -32.8425
22.1035 | -35.3426 | -32.8487
22.1096 | -35.343 | -32.8551
22.1157 | -35.3434 | -32.8616
22.1218 | -35.3439 | -32.8682
22.1279 | -35.3444 | -32.8741
22.134 | -35.3448 | -32.8805
22.1402 | -35.3452 | -32.8868
22.1463 | -35.3456 | -32.8928
22.1524 | -35.3461 | -32.8992
22.1585 | -35.3465 | -32.9057
22.1646 | -35.347 | -32.9118
22.1707 | -35.3474 | -32.9185
22.1768 | -35.3479 | -32.9249
22.1829 | -35.3483 | -32.9313
22.189 | -35.3487 | -32.9378
22.1951 | -35.3492 | -32.9437
22.2012 | -35.3496 | -32.9501
22.2073 | -35.35 | -32.9564
22.2134 | -35.3505 | -32.9627
22.2195 | -35.3509 | -32.9689
22.2256 | -35.3513 | -32.9753
22.2317 | -35.3517 | -32.9817
22.2378 | -35.3521 | -32.988
22.2439 | -35.3526 | -32.9945
22.25 | -35.353 | -33.0011
22.2561 | -35.3534 | -33.0073
22.2622 | -35.3538 | -33.0138
22.2683 | -35.3542 | -33.0198
22.2744 | -35.3547 | -33.0257
22.2806 | -35.3551 | -33.0317
22.2867 | -35.3555 | -33.0379
22.2928 | -35.356 | -33.0443
22.2989 | -35.3564 | -33.0505
22.305 | -35.3568 | -33.0571
22.3111 | -35.3572 | -33.0634
22.3172 | -35.3577 | -33.0697
22.3233 | -35.3581 | -33.0761
22.3294 | -35.3585 | -33.0822
22.3355 | -35.359 | -33.0884
22.3416 | -35.3594 | -33.0946
22.3477 | -35.3599 | -33.1009
22.3538 | -35.3603 | -33.1071
22.3599 | -35.3607 | -33.1134
22.366 | -35.3611 | -33.1196
22.3721 | -35.3615 | -33.1257
22.3782 | -35.3619 | -33.1321
22.3843 | -35.3624 | -33.1385
22.3904 | -35.3628 | -33.1449
22.3965 | -35.3632 | -33.1513
22.4026 | -35.3637 | -33.1572
22.4087 | -35.364 | -33.1636
22.4148 | -35.3645 | -33.1698
22.4209 | -35.3649 | -33.1758
22.4271 | -35.3653 | -33.182
22.4332 | -35.3657 | -33.1882
22.4393 | -35.3662 | -33.1945
22.4454 | -35.3666 | -33.2004
22.4515 | -35.3669 | -33.2065
22.4576 | -35.3673 | -33.2128
22.4637 | -35.3677 | -33.219
22.4698 | -35.3682 | -33.2253
22.4759 | -35.3686 | -33.2315
22.482 | -35.369 | -33.2378
22.4881 | -35.3694 | -33.2438
22.4942 | -35.3699 | -33.25
22.5003 | -35.3703 | -33.2564
22.5064 | -35.3707 | -33.2625
22.5125 | -35.371 | -33.2688
22.5186 | -35.3715 | -33.2749
22.5247 | -35.3719 | -33.2811
22.5308 | -35.3723 | -33.2875
22.5369 | -35.3728 | -33.294
22.543 | -35.3732 | -33.3001
22.5491 | -35.3735 | -33.306
22.5552 | -35.3739 | -33.3122
22.5613 | -35.3744 | -33.3186
22.5675 | -35.3748 | -33.3251
22.5736 | -35.3752 | -33.3314
22.5797 | -35.3756 | -33.3377
22.5858 | -35.376 | -33.3436
22.5919 | -35.3764 | -33.3498
22.598 | -35.3769 | -33.3563
22.6041 | -35.3773 | -33.3624
22.6102 | -35.3777 | -33.3688
22.6163 | -35.3781 | -33.3747
22.6224 | -35.3785 | -33.3809
22.6285 | -35.379 | -33.3872
22.6346 | -35.3794 | -33.3932
22.6407 | -35.3798 | -33.3997
22.6468 | -35.3802 | -33.406
22.6529 | -35.3805 | -33.4122
22.659 | -35.381 | -33.4182
22.6651 | -35.3813 | -33.4246
22.6712 | -35.3817 | -33.4305
22.6773 | -35.3822 | -33.4365
22.6834 | -35.3826 | -33.4429
22.6895 | -35.383 | -33.4491
22.6956 | -35.3834 | -33.4552
22.7017 | -35.3839 | -33.4614
22.7079 | -35.3843 | -33.468
22.714 | -35.3846 | -33.4741
22.7201 | -35.3851 | -33.4804
22.7262 | -35.3855 | -33.4865
22.7323 | -35.3859 | -33.4925
22.7384 | -35.3863 | -33.4986
22.7445 | -35.3867 | -33.5051
22.7506 | -35.3871 | -33.5112
22.7567 | -35.3876 | -33.5173
22.7628 | -35.3879 | -33.5234
22.7689 | -35.3883 | -33.5294
22.775 | -35.3887 | -33.5359
22.7811 | -35.3891 | -33.5422
22.7872 | -35.3895 | -33.5483
22.7933 | -35.39 | -33.5541
22.7994 | -35.3904 | -33.5603
22.8055 | -35.3908 | -33.5667
22.8116 | -35.3913 | -33.5728
22.8177 | -35.3916 | -33.579
22.8238 | -35.392 | -33.5852
22.8299 | -35.3924 | -33.5912
22.836 | -35.3928 | -33.5975
22.8421 | -35.3933 | -33.6037
22.8482 | -35.3937 | -33.6097
22.8544 | -35.3941 | -33.6159
22.8605 | -35.3946 | -33.6222
22.8666 | -35.3949 | -33.6283
22.8727 | -35.3954 | -33.6347
22.8788 | -35.3958 | -33.6408
22.8849 | -35.3962 | -33.647
22.891 | -35.3966 | -33.6531
22.8971 | -35.397 | -33.6591
22.9032 | -35.3974 | -33.665
22.9093 | -35.3978 | -33.6712
22.9154 | -35.3982 | -33.6773
22.9215 | -35.3986 | -33.6835
22.9276 | -35.3991 | -33.6897
22.9337 | -35.3994 | -33.6955
22.9398 | -35.3998 | -33.7018
22.9459 | -35.4002 | -33.7079
22.952 | -35.4006 | -33.7142
22.9581 | -35.401 | -33.7203
22.9642 | -35.4014 | -33.7266
22.9703 | -35.4018 | -33.7326
22.9764 | -35.4022 | -33.7389
22.9825 | -35.4026 | -33.7448
22.9886 | -35.403 | -33.7508
22.9948 | -35.4034 | -33.7569
23.0009 | -35.4038 | -33.7629
23.007 | -35.4042 | -33.769
23.0131 | -35.4046 | -33.7751
23.0192 | -35.405 | -33.781
23.0253 | -35.4055 | -33.7871
23.0314 | -35.4059 | -33.7931
23.0375 | -35.4063 | -33.7992
23.0436 | -35.4068 | -33.805
23.0497 | -35.4072 | -33.8111
23.0558 | -35.4076 | -33.8173
23.0619 | -35.408 | -33.8236
23.068 | -35.4084 | -33.8298
23.0741 | -35.4087 | -33.8359
23.0802 | -35.4091 | -33.8418
23.0863 | -35.4095 | -33.848
23.0924 | -35.4099 | -33.8543
23.0985 | -35.4103 | -33.8605
23.1046 | -35.4107 | -33.8664
23.1107 | -35.4111 | -33.8723
23.1168 | -35.4115 | -33.8784
23.1229 | -35.4119 | -33.8845
23.129 | -35.4124 | -33.8907
23.1351 | -35.4128 | -33.8968
23.1413 | -35.4132 | -33.903
23.1474 | -35.4136 | -33.9086
23.1535 | -35.414 | -33.9147
23.1596 | -35.4144 | -33.9206
23.1657 | -35.4148 | -33.9267
23.1718 | -35.4151 | -33.933
23.1779 | -35.4155 | -33.9393
23.184 | -35.4159 | -33.9448
23.1901 | -35.4163 | -33.9511
23.1962 | -35.4167 | -33.957
23.2023 | -35.4171 | -33.9631
23.2084 | -35.4175 | -33.9692
23.2145 | -35.4179 | -33.9752
23.2206 | -35.4183 | -33.9814
23.2267 | -35.4188 | -33.9874
23.2328 | -35.4191 | -33.9934
23.2389 | -35.4196 | -33.9995
23.245 | -35.42 | -34.0057
23.2511 | -35.4204 | -34.0116
23.2572 | -35.4208 | -34.0178
23.2633 | -35.4212 | -34.0237
23.2694 | -35.4216 | -34.0299
23.2755 | -35.422 | -34.0361
23.2817 | -35.4224 | -34.042
23.2878 | -35.4227 | -34.0481
23.2939 | -35.4232 | -34.0541
23.3 | -35.4236 | -34.0603
23.3061 | -35.424 | -34.0664
23.3122 | -35.4244 | -34.0724
23.3183 | -35.4248 | -34.0784
23.3244 | -35.4252 | -34.0845
23.3305 | -35.4256 | -34.0907
23.3366 | -35.4259 | -34.097
23.3427 | -35.4263 | -34.1028
23.3488 | -35.4267 | -34.109
23.3549 | -35.4271 | -34.115
23.361 | -35.4275 | -34.1208
23.3671 | -35.428 | -34.1271
23.3732 | -35.4284 | -34.1329
23.3793 | -35.4288 | -34.1388
23.3854 | -35.4292 | -34.1451
23.3915 | -35.4295 | -34.1511
23.3976 | -35.4299 | -34.1575
23.4037 | -35.4303 | -34.1634
23.4098 | -35.4307 | -34.1693
23.4159 | -35.4312 | -34.1754
23.422 | -35.4315 | -34.1818
23.4282 | -35.4319 | -34.1876
23.4343 | -35.4323 | -34.1936
23.4404 | -35.4327 | -34.2
23.4465 | -35.4331 | -34.2059
23.4526 | -35.4335 | -34.212
23.4587 | -35.4338 | -34.218
23.4648 | -35.4342 | -34.2238
23.4709 | -35.4346 | -34.23
23.477 | -35.4349 | -34.236
23.4831 | -35.4353 | -34.2419
23.4892 | -35.4357 | -34.248
23.4953 | -35.4361 | -34.2541
23.5014 | -35.4365 | -34.26
23.5075 | -35.4369 | -34.266
23.5136 | -35.4373 | -34.2719
23.5197 | -35.4376 | -34.278
23.5258 | -35.438 | -34.284
23.5319 | -35.4384 | -34.2897
23.538 | -35.4388 | -34.2959
23.5441 | -35.4392 | -34.3018
23.5502 | -35.4396 | -34.308
23.5563 | -35.44 | -34.3136
23.5624 | -35.4403 | -34.3198
23.5686 | -35.4407 | -34.3255
23.5747 | -35.4411 | -34.3313
23.5808 | -35.4414 | -34.3373
23.5869 | -35.4419 | -34.3432
23.593 | -35.4423 | -34.349
23.5991 | -35.4427 | -34.3552
23.6052 | -35.4431 | -34.3609
23.6113 | -35.4435 | -34.3669
23.6174 | -35.4438 | -34.373
23.6235 | -35.4442 | -34.3789
23.6296 | -35.4446 | -34.3851
23.6357 | -35.4449 | -34.3915
23.6418 | -35.4454 | -34.3971
23.6479 | -35.4457 | -34.4031
23.654 | -35.4461 | -34.409
23.6601 | -35.4466 | -34.4149
23.6662 | -35.4469 | -34.4209
23.6723 | -35.4474 | -34.4267
23.6784 | -35.4478 | -34.4326
23.6845 | -35.4482 | -34.4386
23.6906 | -35.4486 | -34.4447
23.6967 | -35.449 | -34.4504
23.7028 | -35.4493 | -34.4564
23.7089 | -35.4497 | -34.4628
23.7151 | -35.4501 | -34.4689
23.7212 | -35.4505 | -34.4749
23.7273 | -35.4508 | -34.481
23.7334 | -35.4512 | -34.4871
23.7395 | -35.4517 | -34.4932
23.7456 | -35.452 | -34.4993
23.7517 | -35.4524 | -34.5052
23.7578 | -35.4528 | -34.511
23.7639 | -35.4532 | -34.5172
23.77 | -35.4535 | -34.5232
23.7761 | -35.454 | -34.5288
23.7822 | -35.4543 | -34.5348
23.7883 | -35.4547 | -34.5404
23.7944 | -35.4551 | -34.5466
23.8005 | -35.4554 | -34.5529
23.8066 | -35.4558 | -34.5588
23.8127 | -35.4562 | -34.565
23.8188 | -35.4566 | -34.5714
23.8249 | -35.457 | -34.5774
23.831 | -35.4573 | -34.5835
23.8371 | -35.4577 | -34.5896
23.8432 | -35.4581 | -34.5956
23.8493 | -35.4585 | -34.6011
23.8555 | -35.4589 | -34.6071
23.8616 | -35.4593 | -34.6133
23.8677 | -35.4597 | -34.6195
23.8738 | -35.46 | -34.6254
23.8799 | -35.4605 | -34.6314
23.886 | -35.4608 | -34.6365
23.8921 | -35.4611 | -34.6425
23.8982 | -35.4615 | -34.6485
23.9043 | -35.4619 | -34.6545
23.9104 | -35.4623 | -34.6606
23.9165 | -35.4626 | -34.6663
23.9226 | -35.463 | -34.6716
23.9287 | -35.4634 | -34.6774
23.9348 | -35.4637 | -34.6835
23.9409 | -35.4641 | -34.6899
23.947 | -35.4644 | -34.6959
23.9531 | -35.4648 | -34.7016
23.9592 | -35.4652 | -34.7075
23.9653 | -35.4656 | -34.7136
23.9714 | -35.466 | -34.7194
23.9775 | -35.4664 | -34.7255
23.9836 | -35.4668 | -34.7317
23.9897 | -35.4672 | -34.7377
23.9958 | -35.4676 | -34.7438
24.002 | -35.468 | -34.7496
24.0081 | -35.4684 | -34.7553
24.0142 | -35.4688 | -34.7612
24.0203 | -35.4691 | -34.7668
24.0264 | -35.4695 | -34.7727
24.0325 | -35.4699 | -34.7786
24.0386 | -35.4703 | -34.7846
24.0447 | -35.4707 | -34.7905
24.0508 | -35.4711 | -34.7966
24.0569 | -35.4715 | -34.8028
24.063 | -35.4718 | -34.8088
24.0691 | -35.4722 | -34.8145
24.0752 | -35.4726 | -34.8205
24.0813 | -35.473 | -34.8264
24.0874 | -35.4735 | -34.8322
24.0935 | -35.4738 | -34.8381
24.0996 | -35.4742 | -34.8437
24.1057 | -35.4746 | -34.8497
24.1118 | -35.4749 | -34.8558
24.1179 | -35.4753 | -34.8617
24.124 | -35.4757 | -34.8673
24.1301 | -35.476 | -34.873
24.1362 | -35.4764 | -34.8789
24.1424 | -35.4768 | -34.8848
24.1485 | -35.4771 | -34.8905
24.1546 | -35.4775 | -34.8966
24.1607 | -35.4779 | -34.9023
24.1668 | -35.4783 | -34.9083
24.1729 | -35.4787 | -34.9142
24.179 | -35.4791 | -34.9201
24.1851 | -35.4795 | -34.9262
24.1912 | -35.4799 | -34.9322
24.1973 | -35.4803 | -34.9382
24.2034 | -35.4807 | -34.9441
24.2095 | -35.4811 | -34.9501
24.2156 | -35.4815 | -34.956
24.2217 | -35.4818 | -34.9619
24.2278 | -35.4822 | -34.9679
24.2339 | -35.4826 | -34.9737
24.24 | -35.483 | -34.9796
24.2461 | -35.4834 | -34.9856
24.2522 | -35.4838 | -34.9915
24.2583 | -35.4842 | -34.9976
24.2644 | -35.4846 | -35.0037
24.2705 | -35.485 | -35.0097
24.2766 | -35.4853 | -35.0153
24.2827 | -35.4857 | -35.0213
24.2889 | -35.4861 | -35.0272
24.295 | -35.4864 | -35.0331
24.3011 | -35.4868 | -35.0389
24.3072 | -35.4872 | -35.0444
24.3133 | -35.4875 | -35.05
24.3194 | -35.4879 | -35.0561
24.3255 | -35.4883 | -35.0624
24.3316 | -35.4886 | -35.0683
24.3377 | -35.489 | -35.0741
24.3438 | -35.4894 | -35.08
24.3499 | -35.4897 | -35.0859
24.356 | -35.4901 | -35.0921
24.3621 | -35.4905 | -35.098
24.3682 | -35.4909 | -35.1037
24.3743 | -35.4913 | -35.11
24.3804 | -35.4916 | -35.1154
24.3865 | -35.492 | -35.1215
24.3926 | -35.4924 | -35.1272
24.3987 | -35.4927 | -35.1331
24.4048 | -35.4931 | -35.139
24.4109 | -35.4935 | -35.145
24.417 | -35.4939 | -35.1504
24.4231 | -35.4943 | -35.1564
24.4293 | -35.4947 | -35.1622
24.4354 | -35.4951 | -35.1679
24.4415 | -35.4954 | -35.1735
24.4476 | -35.4958 | -35.1791
24.4537 | -35.4962 | -35.1849
24.4598 | -35.4966 | -35.1909
24.4659 | -35.497 | -35.1968
24.472 | -35.4974 | -35.2028
24.4781 | -35.4978 | -35.2088
24.4842 | -35.4981 | -35.2146
24.4903 | -35.4984 | -35.2203
24.4964 | -35.4988 | -35.2262
24.5025 | -35.4991 | -35.2321
24.5086 | -35.4996 | -35.2379
24.5147 | -35.5 | -35.2435
24.5208 | -35.5004 | -35.2492
24.5269 | -35.5007 | -35.2548
24.533 | -35.5011 | -35.2604
24.5391 | -35.5014 | -35.2663
24.5452 | -35.5018 | -35.2721
24.5513 | -35.5022 | -35.2782
24.5574 | -35.5026 | -35.2842
24.5635 | -35.503 | -35.2897
24.5696 | -35.5033 | -35.2955
24.5758 | -35.5037 | -35.3013
24.5819 | -35.5041 | -35.3071
24.588 | -35.5045 | -35.3127
24.5941 | -35.5049 | -35.3182
24.6002 | -35.5052 | -35.3239
24.6063 | -35.5056 | -35.3299
24.6124 | -35.506 | -35.3359
24.6185 | -35.5064 | -35.3416
24.6246 | -35.5068 | -35.3472
24.6307 | -35.5071 | -35.3531
24.6368 | -35.5075 | -35.3588
24.6429 | -35.5079 | -35.3644
24.649 | -35.5083 | -35.37
24.6551 | -35.5087 | -35.3761
24.6612 | -35.509 | -35.3813
24.6673 | -35.5094 | -35.3874
24.6734 | -35.5098 | -35.3934
24.6795 | -35.5101 | -35.3995
24.6856 | -35.5105 | -35.4052
24.6917 | -35.5109 | -35.4112
24.6978 | -35.5113 | -35.4168
24.7039 | -35.5117 | -35.4225
24.71 | -35.5121 | -35.4283
24.7162 | -35.5124 | -35.4343
24.7223 | -35.5128 | -35.4401
24.7284 | -35.5132 | -35.4458
24.7345 | -35.5136 | -35.4517
24.7406 | -35.514 | -35.4572
24.7467 | -35.5143 | -35.4633
24.7528 | -35.5146 | -35.4691
24.7589 | -35.515 | -35.4751
24.765 | -35.5154 | -35.4803
24.7711 | -35.5158 | -35.486
24.7772 | -35.5161 | -35.4921
24.7833 | -35.5165 | -35.498
24.7894 | -35.5168 | -35.5039
24.7955 | -35.5172 | -35.5098
24.8016 | -35.5175 | -35.5156
24.8077 | -35.518 | -35.5218
24.8138 | -35.5183 | -35.5277
24.8199 | -35.5187 | -35.5338
24.826 | -35.5191 | -35.5392
24.8321 | -35.5194 | -35.5451
24.8382 | -35.5198 | -35.5505
24.8443 | -35.5202 | -35.5562
24.8504 | -35.5206 | -35.5618
24.8565 | -35.5209 | -35.5676
24.8627 | -35.5213 | -35.5733
24.8688 | -35.5217 | -35.5792
24.8749 | -35.5221 | -35.5848
24.881 | -35.5224 | -35.5908
24.8871 | -35.5228 | -35.5965
24.8932 | -35.5232 | -35.602
24.8993 | -35.5236 | -35.6075
24.9054 | -35.524 | -35.6133
24.9115 | -35.5244 | -35.6191
24.9176 | -35.5247 | -35.625
24.9237 | -35.5251 | -35.6308
24.9298 | -35.5255 | -35.6365
24.9359 | -35.5259 | -35.6423
24.942 | -35.5262 | -35.6483
24.9481 | -35.5265 | -35.6542
24.9542 | -35.5269 | -35.6602
24.9603 | -35.5272 | -35.6659
24.9664 | -35.5276 | -35.6719
24.9725 | -35.528 | -35.6778
24.9786 | -35.5284 | -35.6834
24.9847 | -35.5287 | -35.689
24.9908 | -35.5291 | -35.6947
24.9969 | -35.5294 | -35.7005
25.0031 | -35.5298 | -35.7064
25.0092 | -35.5301 | -35.712
25.0153 | -35.5305 | -35.7177
25.0214 | -35.5309 | -35.7233
25.0275 | -35.5313 | -35.7288
25.0336 | -35.5317 | -35.7345
25.0397 | -35.5321 | -35.7406
25.0458 | -35.5324 | -35.7463
25.0519 | -35.5328 | -35.7525
25.058 | -35.5332 | -35.758
25.0641 | -35.5336 | -35.7639
25.0702 | -35.5339 | -35.7695
25.0763 | -35.5343 | -35.7754
25.0824 | -35.5347 | -35.7809
25.0885 | -35.5351 | -35.7864
25.0946 | -35.5354 | -35.7922
25.1007 | -35.5358 | -35.7984
25.1068 | -35.5362 | -35.8039
25.1129 | -35.5365 | -35.8096
25.119 | -35.5368 | -35.8148
25.1251 | -35.5372 | -35.821
25.1312 | -35.5376 | -35.8269
25.1373 | -35.538 | -35.8325
25.1435 | -35.5383 | -35.8385
25.1496 | -35.5387 | -35.8442
25.1557 | -35.539 | -35.8502
25.1618 | -35.5394 | -35.8561
25.1679 | -35.5398 | -35.8619
25.174 | -35.5401 | -35.8677
25.1801 | -35.5404 | -35.8736
25.1862 | -35.5409 | -35.8794
25.1923 | -35.5412 | -35.8854
25.1984 | -35.5416 | -35.8914
25.2045 | -35.5421 | -35.8972
25.2106 | -35.5424 | -35.9031
25.2167 | -35.5428 | -35.9088
25.2228 | -35.5432 | -35.9142
25.2289 | -35.5435 | -35.9197
25.235 | -35.5439 | -35.9254
25.2411 | -35.5443 | -35.931
25.2472 | -35.5446 | -35.9369
25.2533 | -35.545 | -35.943
25.2594 | -35.5454 | -35.9484
25.2655 | -35.5458 | -35.9542
25.2716 | -35.5462 | -35.96
25.2777 | -35.5465 | -35.9659
25.2838 | -35.5469 | -35.9719
25.29 | -35.5473 | -35.9777
25.2961 | -35.5476 | -35.9838
25.3022 | -35.548 | -35.9895
25.3083 | -35.5483 | -35.9951
25.3144 | -35.5486 | -36.0005
25.3205 | -35.549 | -36.0061
25.3266 | -35.5493 | -36.0121
25.3327 | -35.5497 | -36.0183
25.3388 | -35.55 | -36.0239
25.3449 | -35.5504 | -36.0298
25.351 | -35.5508 | -36.0356
25.3571 | -35.5511 | -36.0416
25.3632 | -35.5515 | -36.0474
25.3693 | -35.5518 | -36.0532
25.3754 | -35.5521 | -36.0593
25.3815 | -35.5524 | -36.0651
25.3876 | -35.5528 | -36.0709
25.3937 | -35.5531 | -36.0766
25.3998 | -35.5535 | -36.0821
25.4059 | -35.5539 | -36.0881
25.412 | -35.5542 | -36.0939
25.4181 | -35.5546 | -36.0997
25.4242 | -35.555 | -36.1056
25.4304 | -35.5554 | -36.1114
25.4365 | -35.5558 | -36.1172
25.4426 | -35.5561 | -36.1232
25.4487 | -35.5564 | -36.1292
25.4548 | -35.5568 | -36.1349
25.4609 | -35.5571 | -36.1408
25.467 | -35.5576 | -36.1461
25.4731 | -35.558 | -36.152
25.4792 | -35.5583 | -36.1578
25.4853 | -35.5587 | -36.1636
25.4914 | -35.5591 | -36.1696
25.4975 | -35.5595 | -36.175
25.5036 | -35.5599 | -36.1803
25.5097 | -35.5602 | -36.1861
25.5158 | -35.5605 | -36.1917
25.5219 | -35.5609 | -36.1975
25.528 | -35.5612 | -36.2033
25.5341 | -35.5616 | -36.2087
25.5402 | -35.5619 | -36.2146
25.5463 | -35.5623 | -36.2203
25.5524 | -35.5627 | -36.226
25.5585 | -35.563 | -36.2315
25.5646 | -35.5634 | -36.2373
25.5707 | -35.5638 | -36.243
25.5769 | -35.5642 | -36.2493
25.583 | -35.5646 | -36.2551
25.5891 | -35.565 | -36.261
25.5952 | -35.5653 | -36.2667
25.6013 | -35.5657 | -36.2725
25.6074 | -35.5661 | -36.2778
25.6135 | -35.5666 | -36.2836
25.6196 | -35.567 | -36.2897
25.6257 | -35.5673 | -36.2952
25.6318 | -35.5677 | -36.3009
25.6379 | -35.568 | -36.3067
25.644 | -35.5684 | -36.312
25.6501 | -35.5687 | -36.3174
25.6562 | -35.5691 | -36.3233
25.6623 | -35.5694 | -36.3291
25.6684 | -35.5698 | -36.335
25.6745 | -35.5701 | -36.3409
25.6806 | -35.5705 | -36.3468
25.6867 | -35.5708 | -36.3528
25.6928 | -35.5712 | -36.3585
25.6989 | -35.5716 | -36.3638
25.705 | -35.5719 | -36.3698
25.7111 | -35.5722 | -36.3754
25.7173 | -35.5725 | -36.3811
25.7234 | -35.5729 | -36.3867
25.7295 | -35.5732 | -36.3925
25.7356 | -35.5735 | -36.3984
25.7417 | -35.5739 | -36.404
25.7478 | -35.5743 | -36.4097
25.7539 | -35.5746 | -36.4151
25.76 | -35.575 | -36.4209
25.7661 | -35.5754 | -36.4266
25.7722 | -35.5757 | -36.4322
25.7783 | -35.5761 | -36.438
25.7844 | -35.5765 | -36.4436
25.7905 | -35.5768 | -36.4494
25.7966 | -35.5772 | -36.4552
25.8027 | -35.5775 | -36.461
25.8088 | -35.5779 | -36.4668
25.8149 | -35.5783 | -36.4723
25.821 | -35.5786 | -36.4778
25.8271 | -35.579 | -36.4837
25.8332 | -35.5793 | -36.4898
25.8393 | -35.5797 | -36.4953
25.8454 | -35.58 | -36.5008
25.8515 | -35.5805 | -36.5064
25.8576 | -35.5808 | -36.5124
25.8638 | -35.5812 | -36.5178
25.8699 | -35.5815 | -36.5235
25.876 | -35.5819 | -36.5294
25.8821 | -35.5822 | -36.5354
25.8882 | -35.5825 | -36.5412
25.8943 | -35.5829 | -36.5471
25.9004 | -35.5832 | -36.5527
25.9065 | -35.5836 | -36.5587
25.9126 | -35.5839 | -36.5644
25.9187 | -35.5843 | -36.5703
25.9248 | -35.5847 | -36.5762
25.9309 | -35.585 | -36.5819
25.937 | -35.5854 | -36.5877
25.9431 | -35.5858 | -36.5932
25.9492 | -35.5862 | -36.5987
25.9553 | -35.5865 | -36.6042
25.9614 | -35.5869 | -36.6097
25.9675 | -35.5871 | -36.6152
25.9736 | -35.5875 | -36.6209
25.9797 | -35.5879 | -36.6265
25.9858 | -35.5882 | -36.6321
25.9919 | -35.5886 | -36.6378
25.998 | -35.589 | -36.644
26.0042 | -35.5893 | -36.6497
26.0103 | -35.5896 | -36.6554
26.0164 | -35.5899 | -36.661
26.0225 | -35.5903 | -36.6669
26.0286 | -35.5906 | -36.6726
26.0347 | -35.591 | -36.678
26.0408 | -35.5913 | -36.6835
26.0469 | -35.5917 | -36.6888
26.053 | -35.592 | -36.6945
26.0591 | -35.5924 | -36.7003
26.0652 | -35.5927 | -36.7063
26.0713 | -35.5931 | -36.7119
26.0774 | -35.5934 | -36.7177
26.0835 | -35.5938 | -36.7233
26.0896 | -35.5941 | -36.7289
26.0957 | -35.5945 | -36.7345
26.1018 | -35.5949 | -36.7403
26.1079 | -35.5953 | -36.7459
26.114 | -35.5956 | -36.7517
26.1201 | -35.5959 | -36.7574
26.1262 | -35.5963 | -36.7631
26.1323 | -35.5966 | -36.7685
26.1384 | -35.597 | -36.7744
26.1445 | -35.5974 | -36.78
26.1507 | -35.5978 | -36.7858
26.1568 | -35.5981 | -36.7915
26.1629 | -35.5984 | -36.7973
26.169 | -35.5988 | -36.8027
26.1751 | -35.5991 | -36.8081
26.1812 | -35.5994 | -36.8139
26.1873 | -35.5998 | -36.8194
26.1934 | -35.6002 | -36.8249
26.1995 | -35.6005 | -36.8304
26.2056 | -35.6009 | -36.8353
26.2117 | -35.6012 | -36.8413
26.2178 | -35.6016 | -36.8469
26.2239 | -35.6019 | -36.8522
26.23 | -35.6023 | -36.8581
26.2361 | -35.6026 | -36.8635
26.2422 | -35.603 | -36.8697
26.2483 | -35.6033 | -36.8752
26.2544 | -35.6038 | -36.8807
26.2605 | -35.6041 | -36.8864
26.2666 | -35.6044 | -36.8919
26.2727 | -35.6047 | -36.8974
26.2788 | -35.6051 | -36.9029
26.2849 | -35.6055 | -36.9085
26.2911 | -35.6058 | -36.9144
26.2972 | -35.6061 | -36.9201
26.3033 | -35.6065 | -36.9257
26.3094 | -35.6068 | -36.9313
26.3155 | -35.6071 | -36.9366
26.3216 | -35.6074 | -36.9426
26.3277 | -35.6078 | -36.948
26.3338 | -35.6081 | -36.9539
26.3399 | -35.6084 | -36.9592
26.346 | -35.6087 | -36.9646
26.3521 | -35.6091 | -36.9705
26.3582 | -35.6094 | -36.9766
26.3643 | -35.6097 | -36.9823
26.3704 | -35.6101 | -36.9883
26.3765 | -35.6104 | -36.9941
26.3826 | -35.6108 | -36.9999
26.3887 | -35.6111 | -37.0056
26.3948 | -35.6115 | -37.0108
26.4009 | -35.6118 | -37.0162
26.407 | -35.6122 | -37.0217
26.4131 | -35.6125 | -37.0274
26.4192 | -35.6129 | -37.0337
26.4253 | -35.6132 | -37.039
26.4314 | -35.6136 | -37.0442
26.4376 | -35.614 | -37.0498
26.4437 | -35.6143 | -37.0555
26.4498 | -35.6146 | -37.0614
26.4559 | -35.615 | -37.0663
26.462 | -35.6153 | -37.0724
26.4681 | -35.6157 | -37.0784
26.4742 | -35.6161 | -37.084
26.4803 | -35.6164 | -37.0896
26.4864 | -35.6168 | -37.095
26.4925 | -35.6171 | -37.1008
26.4986 | -35.6174 | -37.1062
26.5047 | -35.6178 | -37.1128
26.5108 | -35.6182 | -37.1181
26.5169 | -35.6185 | -37.1234
26.523 | -35.6188 | -37.1291
26.5291 | -35.6192 | -37.1344
26.5352 | -35.6196 | -37.1403
26.5413 | -35.6199 | -37.146
26.5474 | -35.6203 | -37.1517
26.5535 | -35.6206 | -37.157
26.5596 | -35.6208 | -37.1623
26.5657 | -35.6213 | -37.1679
26.5718 | -35.6216 | -37.1732
26.578 | -35.622 | -37.1789
26.5841 | -35.6223 | -37.1847
26.5902 | -35.6226 | -37.1904
26.5963 | -35.623 | -37.1957
26.6024 | -35.6233 | -37.2016
26.6085 | -35.6237 | -37.2071
26.6146 | -35.624 | -37.2124
26.6207 | -35.6244 | -37.2184
26.6268 | -35.6248 | -37.2243
26.6329 | -35.6251 | -37.2296
26.639 | -35.6254 | -37.2354
26.6451 | -35.6256 | -37.241
26.6512 | -35.626 | -37.2464
26.6573 | -35.6264 | -37.2524
26.6634 | -35.6267 | -37.2578
26.6695 | -35.6271 | -37.2638
26.6756 | -35.6274 | -37.2698
26.6817 | -35.6278 | -37.2759
26.6878 | -35.628 | -37.2812
26.6939 | -35.6284 | -37.2867
26.7 | -35.6288 | -37.2916
26.7061 | -35.6291 | -37.2974
26.7122 | -35.6295 | -37.3032
26.7183 | -35.6298 | -37.3088
26.7245 | -35.6302 | -37.3141
26.7306 | -35.6305 | -37.3195
26.7367 | -35.6308 | -37.3246
26.7428 | -35.6312 | -37.3299
26.7489 | -35.6317 | -37.3357
26.755 | -35.6319 | -37.3412
26.7611 | -35.6322 | -37.3466
26.7672 | -35.6326 | -37.3526
26.7733 | -35.6331 | -37.358
26.7794 | -35.6334 | -37.3637
26.7855 | -35.6337 | -37.3696
26.7916 | -35.6341 | -37.3755
26.7977 | -35.6345 | -37.3805
26.8038 | -35.6348 | -37.3863
26.8099 | -35.6351 | -37.3921
26.816 | -35.6355 | -37.3975
26.8221 | -35.6358 | -37.4028
26.8282 | -35.6362 | -37.4085
26.8343 | -35.6366 | -37.4142
26.8404 | -35.637 | -37.4202
26.8465 | -35.6372 | -37.4259
26.8526 | -35.6376 | -37.432
26.8587 | -35.6379 | -37.4376
26.8649 | -35.6383 | -37.4434
26.871 | -35.6386 | -37.4487
26.8771 | -35.639 | -37.4546
26.8832 | -35.6393 | -37.4598
26.8893 | -35.6397 | -37.4658
26.8954 | -35.64 | -37.4711
26.9015 | -35.6403 | -37.4767
26.9076 | -35.6406 | -37.4832
26.9137 | -35.641 | -37.4888
26.9198 | -35.6414 | -37.4942
26.9259 | -35.6417 | -37.499
26.932 | -35.642 | -37.5042
26.9381 | -35.6423 | -37.5104
26.9442 | -35.6427 | -37.5158
26.9503 | -35.643 | -37.5222
26.9564 | -35.6434 | -37.5278
26.9625 | -35.6438 | -37.5331
26.9686 | -35.6442 | -37.5384
26.9747 | -35.6445 | -37.5445
26.9808 | -35.6449 | -37.55
26.9869 | -35.6453 | -37.5551
26.993 | -35.6456 | -37.5601
26.9991 | -35.646 | -37.5659
27.0052 | -35.6463 | -37.5716
27.0114 | -35.6466 | -37.5769
27.0175 | -35.6471 | -37.5818
27.0236 | -35.6474 | -37.5872
27.0297 | -35.6477 | -37.5926
27.0358 | -35.648 | -37.5981
27.0419 | -35.6484 | -37.604
27.048 | -35.6487 | -37.6089
27.0541 | -35.6489 | -37.6138
27.0602 | -35.6492 | -37.6199
27.0663 | -35.6495 | -37.6257
27.0724 | -35.6499 | -37.6316
27.0785 | -35.6503 | -37.6369
27.0846 | -35.6506 | -37.6426
27.0907 | -35.6509 | -37.6486
27.0968 | -35.6514 | -37.6539
27.1029 | -35.6518 | -37.6592
27.109 | -35.6521 | -37.665
27.1151 | -35.6525 | -37.6705
27.1212 | -35.6528 | -37.6758
27.1273 | -35.6531 | -37.6817
27.1334 | -35.6535 | -37.6878
27.1395 | -35.6538 | -37.6934
27.1456 | -35.6542 | -37.6986
27.1518 | -35.6545 | -37.7046
27.1579 | -35.6549 | -37.7102
27.164 | -35.6552 | -37.7161
27.1701 | -35.6555 | -37.7218
27.1762 | -35.6558 | -37.7275
27.1823 | -35.6562 | -37.733
27.1884 | -35.6565 | -37.7386
27.1945 | -35.6567 | -37.743
27.2006 | -35.657 | -37.7481
27.2067 | -35.6574 | -37.7533
27.2128 | -35.6578 | -37.7589
27.2189 | -35.658 | -37.764
27.225 | -35.6584 | -37.7695
27.2311 | -35.6587 | -37.775
27.2372 | -35.659 | -37.7805
27.2433 | -35.6593 | -37.7856
27.2494 | -35.6597 | -37.7913
27.2555 | -35.66 | -37.7974
27.2616 | -35.6602 | -37.8031
27.2677 | -35.6607 | -37.8087
27.2738 | -35.661 | -37.814
27.2799 | -35.6613 | -37.8196
27.286 | -35.6616 | -37.8249
27.2921 | -35.6619 | -37.8304
27.2983 | -35.6622 | -37.8358
27.3044 | -35.6625 | -37.8414
27.3105 | -35.6628 | -37.8475
27.3166 | -35.6631 | -37.8526
27.3227 | -35.6635 | -37.8584
27.3288 | -35.6638 | -37.8639
27.3349 | -35.664 | -37.8693
27.341 | -35.6644 | -37.8749
27.3471 | -35.6647 | -37.8803
27.3532 | -35.6651 | -37.8851
27.3593 | -35.6656 | -37.8908
27.3654 | -35.6659 | -37.8962
27.3715 | -35.6663 | -37.9015
27.3776 | -35.6666 | -37.9071
27.3837 | -35.6669 | -37.9122
27.3898 | -35.6672 | -37.918
27.3959 | -35.6676 | -37.9232
27.402 | -35.668 | -37.9287
27.4081 | -35.6684 | -37.9336
27.4142 | -35.6688 | -37.9391
27.4203 | -35.6692 | -37.9444
27.4264 | -35.6694 | -37.9504
27.4325 | -35.6698 | -37.9558
27.4387 | -35.6701 | -37.9607
27.4448 | -35.6705 | -37.9662
27.4509 | -35.6708 | -37.9714
27.457 | -35.6713 | -37.9766
27.4631 | -35.6715 | -37.9822
27.4692 | -35.6718 | -37.9876
27.4753 | -35.6721 | -37.9929
27.4814 | -35.6724 | -37.9986
27.4875 | -35.6728 | -38.0046
27.4936 | -35.6731 | -38.0098
27.4997 | -35.6734 | -38.0151
27.5058 | -35.6737 | -38.0206
27.5119 | -35.674 | -38.0263
27.518 | -35.6743 | -38.0317
27.5241 | -35.6746 | -38.0375
27.5302 | -35.6749 | -38.0432
27.5363 | -35.6752 | -38.0484
27.5424 | -35.6756 | -38.0538
27.5485 | -35.6759 | -38.0597
27.5546 | -35.6762 | -38.0651
27.5607 | -35.6766 | -38.0705
27.5668 | -35.6769 | -38.076
27.5729 | -35.6772 | -38.0818
27.5791 | -35.6775 | -38.087
27.5852 | -35.6778 | -38.092
27.5913 | -35.6781 | -38.0975
27.5974 | -35.6785 | -38.103
27.6035 | -35.6787 | -38.1081
27.6096 | -35.679 | -38.1131
27.6157 | -35.6793 | -38.119
27.6218 | -35.6796 | -38.1247
27.6279 | -35.6799 | -38.1302
27.634 | -35.6803 | -38.1362
27.6401 | -35.6806 | -38.1412
27.6462 | -35.6809 | -38.1465
27.6523 | -35.6812 | -38.151
27.6584 | -35.6815 | -38.1566
27.6645 | -35.6818 | -38.162
27.6706 | -35.6821 | -38.167
27.6767 | -35.6825 | -38.1731
27.6828 | -35.6828 | -38.178
27.6889 | -35.6831 | -38.1837
27.695 | -35.6834 | -38.1888
27.7011 | -35.6837 | -38.1941
27.7072 | -35.6841 | -38.1997
27.7133 | -35.6844 | -38.2053
27.7194 | -35.6847 | -38.2103
27.7256 | -35.6851 | -38.2157
27.7317 | -35.6854 | -38.2212
27.7378 | -35.6857 | -38.2264
27.7439 | -35.686 | -38.2322
27.75 | -35.6863 | -38.2376
27.7561 | -35.6865 | -38.2432
27.7622 | -35.6869 | -38.2487
27.7683 | -35.6872 | -38.2538
27.7744 | -35.6875 | -38.2598
27.7805 | -35.6877 | -38.265
27.7866 | -35.6881 | -38.2708
27.7927 | -35.6884 | -38.276
27.7988 | -35.6887 | -38.2816
27.8049 | -35.6889 | -38.2873
27.811 | -35.6893 | -38.2922
27.8171 | -35.6896 | -38.2978
27.8232 | -35.6899 | -38.3029
27.8293 | -35.6902 | -38.3085
27.8354 | -35.6905 | -38.314
27.8415 | -35.6909 | -38.3193
27.8476 | -35.6912 | -38.3252
27.8537 | -35.6915 | -38.3309
27.8598 | -35.6918 | -38.3361
27.866 | -35.6921 | -38.3411
27.8721 | -35.6924 | -38.3468
27.8782 | -35.6928 | -38.352
27.8843 | -35.6932 | -38.3574
27.8904 | -35.6934 | -38.3629
27.8965 | -35.6937 | -38.3688
27.9026 | -35.6939 | -38.374
27.9087 | -35.6943 | -38.3795
27.9148 | -35.6946 | -38.3843
27.9209 | -35.6949 | -38.3898
27.927 | -35.6951 | -38.3953
27.9331 | -35.6956 | -38.4007
27.9392 | -35.6959 | -38.4065
27.9453 | -35.6963 | -38.4116
27.9514 | -35.6966 | -38.4172
27.9575 | -35.6969 | -38.4229
27.9636 | -35.6972 | -38.4288
27.9697 | -35.6975 | -38.4333
27.9758 | -35.6978 | -38.4386
27.9819 | -35.6981 | -38.4442
27.988 | -35.6984 | -38.4495
27.9941 | -35.6986 | -38.4553
28.0002 | -35.6988 | -38.4606
28.0063 | -35.699 | -38.4656
28.0125 | -35.6993 | -38.4711
28.0186 | -35.6997 | -38.4767
28.0247 | -35.7001 | -38.4825
28.0308 | -35.7004 | -38.4878
28.0369 | -35.7008 | -38.4931
28.043 | -35.7011 | -38.4987
28.0491 | -35.7014 | -38.5042
28.0552 | -35.7016 | -38.51
28.0613 | -35.7019 | -38.5157
28.0674 | -35.7022 | -38.5207
28.0735 | -35.7026 | -38.5267

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.3555 | -42.6796
21.4908 | -35.5171 | -35.3355
16.4948 | -34.9039 | -27.8788
12.6602 | -34.5477 | -21.5456
9.7171 | -34.3307 | -16.592
7.4582 | -34.1923 | -12.5819
5.7244 | -34.0845 | -9.3378
4.3936 | -34.0196 | -6.7166
3.3722 | -34.0055 | -4.6541
2.5883 | -34.0178 | -3.1505
1.9866 | -34.0057 | -1.8417
1.5248 | -34.0048 | -0.84245
1.1703 | -34.0133 | -0.18734
1 | -34.0199 | 0.13406
0.89822 | -34.008 | 0.40631
0.68941 | -33.9881 | 0.98917
0.52919 | -33.9985 | 1.4421
0.40612 | -34.0034 | 2.0467
0.31171 | -34.0042 | 2.4978
0.23928 | -34.0106 | 3.2191
0.18358 | -34.017 | 3.6298
0.14091 | -34.0335 | 4.5008
0.10819 | -34.0359 | 5.6624
0.083017 | -34.0608 | 7.1322
0.063705 | -34.1161 | 9.0028
0.048923 | -34.1867 | 11.4963
0.037479 | -34.2872 | 14.6946
0.028801 | -34.4704 | 18.692
0.022078 | -34.7834 | 23.964
0.017023 | -35.2589 | 30.3254
0.013018 | -36.0565 | 37.9754
0.010014 | -37.1647 | 45.8686

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003711 | -11.2162
0.007423 | -34.5682
0.011134 | -37.6367
0.014846 | -40.7632
0.018557 | -42.6799
0.022269 | -44.6304
0.02598 | -46.6041
0.029691 | -49.3392
0.033403 | -50.369
0.037114 | -51.2428
0.040826 | -51.6257
0.044537 | -53.9141
0.048249 | -54.218
0.05196 | -54.1962
0.055672 | -54.9691
0.059383 | -55.439
0.063094 | -56.688
0.066806 | -58.7601
0.070517 | -59.6939
0.074229 | -59.5103
0.07794 | -58.8185
0.081652 | -59.0231
0.085363 | -61.9258
0.089074 | -62.0091
0.092786 | -62.037
0.096497 | -62.6907
0.10021 | -62.4929
0.10392 | -61.9698
0.10763 | -63.4124
0.11134 | -63.8831
0.11506 | -64.5607
0.11877 | -64.8118
0.12248 | -64.9159
0.12619 | -64.2824
0.1299 | -64.9744
0.13361 | -65.8154
0.13732 | -66.0865
0.14103 | -66.9555
0.14475 | -67.0777
0.14846 | -67.1934
0.15217 | -66.5267
0.15588 | -66.6698
0.15959 | -67.8984
0.1633 | -69.2264
0.16701 | -68.7098
0.17073 | -69.2907
0.17444 | -69.0515
0.17815 | -69.6434
0.18186 | -70.1752
0.18557 | -70.131
0.18928 | -70.9531
0.19299 | -70.5415
0.19671 | -70.6907
0.20042 | -70.6726
0.20413 | -70.4353
0.20784 | -71.6704
0.21155 | -70.9526
0.21526 | -70.9068
0.21898 | -71.5512
0.22269 | -71.3183
0.2264 | -72.2676
0.23011 | -71.5585
0.23382 | -72.2595
0.23753 | -74.6535
0.24124 | -74.3334
0.24496 | -73.7049
0.24867 | -73.6074
0.25238 | -74.1071
0.25609 | -73.3119
0.2598 | -73.1338
0.26351 | -73.4907
0.26722 | -74.0743
0.27093 | -74.5871
0.27465 | -74.7951
0.27836 | -74.4489
0.28207 | -74.2804
0.28578 | -73.9727
0.28949 | -74.9668
0.2932 | -76.0219
0.29691 | -75.013
0.30063 | -75.1352
0.30434 | -74.7704
0.30805 | -74.6607
0.31176 | -75.3599
0.31547 | -76.4215
0.31918 | -76.3444
0.32289 | -77.2428
0.32661 | -77.5894
0.33032 | -77.27
0.33403 | -76.8128
0.33774 | -76.6854
0.34145 | -77.1429
0.34516 | -77.1184
0.34887 | -76.4381
0.35259 | -77.4258
0.3563 | -78.8425
0.36001 | -77.8247
0.36372 | -77.7627
0.36743 | -77.873
0.37114 | -77.0704
0.37485 | -77.1795
0.37857 | -77.2959
0.38228 | -77.2733
0.38599 | -76.6937
0.3897 | -76.8499
0.39341 | -77.8656
0.39712 | -79.4
0.40083 | -80.0366
0.40455 | -78.903
0.40826 | -78.1544
0.41197 | -79.113
0.41568 | -78.9932
0.41939 | -79.0739
0.4231 | -80.0545
0.42681 | -80.2028
0.43053 | -80.3744
0.43424 | -79.5653
0.43795 | -79.6896
0.44166 | -80.1348
0.44537 | -80.2092
0.44908 | -80.2859
0.4528 | -80.4922
0.45651 | -81.436
0.46022 | -80.3478
0.46393 | -79.5587
0.46764 | -80.2085
0.47135 | -80.8053
0.47506 | -80.144
0.47878 | -80.7585
0.48249 | -80.3689
0.4862 | -79.8446
0.48991 | -80.2951
0.49362 | -80.8111
0.49733 | -81.4835
0.50104 | -82.0109
0.50475 | -81.9076
0.50847 | -81.705
0.51218 | -81.3526
0.51589 | -81.6366
0.5196 | -81.6511
0.52331 | -80.7203
0.52702 | -81.2228
0.53073 | -81.8082
0.53445 | -81.7676
0.53816 | -81.263
0.54187 | -80.7475
0.54558 | -80.985
0.54929 | -81.4396
0.553 | -81.992
0.55671 | -82.1311
0.56043 | -82.01
0.56414 | -82.0553
0.56785 | -81.966
0.57156 | -81.5475
0.57527 | -81.2565
0.57898 | -81.5169
0.58269 | -82.2175
0.58641 | -82.6902
0.59012 | -82.5749
0.59383 | -82.6701
0.59754 | -83.1686
0.60125 | -83.0965
0.60496 | -82.7318
0.60867 | -82.8126
0.61239 | -83.1061
0.6161 | -83.3949
0.61981 | -83.8295
0.62352 | -84.0931
0.62723 | -83.7292
0.63094 | -83.2135
0.63466 | -83.3966
0.63837 | -84.0029
0.64208 | -84.3222
0.64579 | -84.3883
0.6495 | -84.4029
0.65321 | -84.313
0.65692 | -83.7555
0.66064 | -83.0727
0.66435 | -82.6986
0.66806 | -82.7684
0.67177 | -83.3232
0.67548 | -83.8946
0.67919 | -83.9375
0.6829 | -83.9184
0.68662 | -84.0157
0.69033 | -84.0897
0.69404 | -84.134
0.69775 | -84.0373
0.70146 | -83.7933
0.70517 | -83.855
0.70888 | -84.0537
0.7126 | -84.1057
0.71631 | -84.4408
0.72002 | -84.9772
0.72373 | -84.8721
0.72744 | -84.5811
0.73115 | -84.7284
0.73486 | -84.4495
0.73858 | -84.0933
0.74229 | -84.3857
0.746 | -85.1476
0.74971 | -85.6103
0.75342 | -85.3445
0.75713 | -84.9407
0.76084 | -85.1302
0.76456 | -84.9983
0.76827 | -84.8539
0.77198 | -85.4278
0.77569 | -85.9123
0.7794 | -85.4977
0.78311 | -84.9898
0.78682 | -84.8415
0.79054 | -84.8709
0.79425 | -85.0158
0.79796 | -85.4597
0.80167 | -85.7161
0.80538 | -85.2663
0.80909 | -84.5941
0.8128 | -84.0436
0.81652 | -84.3426
0.82023 | -85.4431
0.82394 | -86.317
0.82765 | -86.7605
0.83136 | -86.8185
0.83507 | -86.3292
0.83879 | -85.4544
0.8425 | -85.1681
0.84621 | -85.5493
0.84992 | -85.8134
0.85363 | -85.603
0.85734 | -85.3099
0.86105 | -85.2648
0.86477 | -85.5636
0.86848 | -85.7295
0.87219 | -85.6353
0.8759 | -85.7499
0.87961 | -86.3183
0.88332 | -86.6246
0.88703 | -85.9633
0.89075 | -85.4605
0.89446 | -85.3998
0.89817 | -85.526
0.90188 | -85.6749
0.90559 | -86.1334
0.9093 | -86.572
0.91301 | -86.8414
0.91673 | -86.7224
0.92044 | -86.4192
0.92415 | -86.4015
0.92786 | -86.459
0.93157 | -86.4343
0.93528 | -86.2354
0.93899 | -86.0273
0.94271 | -86.3289
0.94642 | -86.76
0.95013 | -87.2078
0.95384 | -88.0085
0.95755 | -88.4512
0.96126 | -88.3349
0.96497 | -87.9064
0.96869 | -87.5336
0.9724 | -87.7407
0.97611 | -88.0469
0.97982 | -88.2743
0.98353 | -88.2805
0.98724 | -88.1224
0.99095 | -87.8391
0.99467 | -87.5936
0.99838 | -87.2794
1.0021 | -86.9623
1.0058 | -87.2904
1.0095 | -87.5778
1.0132 | -87.4095
1.0169 | -87.2946
1.0206 | -87.389
1.0244 | -87.3411
1.0281 | -87.1559
1.0318 | -87.1936
1.0355 | -87.4275
1.0392 | -87.8322
1.0429 | -88.184
1.0466 | -88.0712
1.0503 | -87.8573
1.054 | -87.8662
1.0578 | -87.985
1.0615 | -88.1409
1.0652 | -88.3378
1.0689 | -88.2514
1.0726 | -88.142
1.0763 | -88.1038
1.08 | -88.1806
1.0837 | -88.2299
1.0875 | -88.2484
1.0912 | -88.3213
1.0949 | -88.344
1.0986 | -88.1606
1.1023 | -87.935
1.106 | -87.9346
1.1097 | -88.1586
1.1134 | -88.4485
1.1171 | -88.6798
1.1209 | -88.8693
1.1246 | -89.1452
1.1283 | -89.3541
1.132 | -89.311
1.1357 | -89.1801
1.1394 | -89.0809
1.1431 | -88.9778
1.1468 | -88.9551
1.1505 | -89.2038
1.1543 | -89.7009
1.158 | -90.0089
1.1617 | -89.7992
1.1654 | -89.3822
1.1691 | -89.0951
1.1728 | -88.9076
1.1765 | -88.6919
1.1802 | -88.4259
1.1839 | -88.1732
1.1877 | -88.1087
1.1914 | -88.3339
1.1951 | -88.683
1.1988 | -88.8958
1.2025 | -88.9633
1.2062 | -89.1069
1.2099 | -89.4075
1.2136 | -89.583
1.2174 | -89.471
1.2211 | -89.3367
1.2248 | -89.4046
1.2285 | -89.594
1.2322 | -89.6331
1.2359 | -89.4222
1.2396 | -89.1937
1.2433 | -89.1057
1.247 | -89.0834
1.2508 | -89.0736
1.2545 | -89.2226
1.2582 | -89.5979
1.2619 | -89.856
1.2656 | -89.6499
1.2693 | -89.2009
1.273 | -88.8667
1.2767 | -88.7037
1.2804 | -88.6653
1.2842 | -88.7587
1.2879 | -88.9787
1.2916 | -89.2252
1.2953 | -89.3842
1.299 | -89.5172
1.3027 | -89.7813
1.3064 | -90.1411
1.3101 | -90.3164
1.3138 | -90.0355
1.3176 | -89.6082
1.3213 | -89.3809
1.325 | -89.3086
1.3287 | -89.3505
1.3324 | -89.5543
1.3361 | -89.8995
1.3398 | -90.2813
1.3435 | -90.5067
1.3473 | -90.495
1.351 | -90.3459
1.3547 | -90.2175
1.3584 | -90.3047
1.3621 | -90.5963
1.3658 | -90.7121
1.3695 | -90.4023
1.3732 | -90.0128
1.3769 | -89.9484
1.3807 | -90.1637
1.3844 | -90.1902
1.3881 | -89.8833
1.3918 | -89.7233
1.3955 | -89.9882
1.3992 | -90.3528
1.4029 | -90.5361
1.4066 | -90.6588
1.4103 | -90.666
1.4141 | -90.536
1.4178 | -90.4833
1.4215 | -90.5629
1.4252 | -90.6998
1.4289 | -90.7872
1.4326 | -90.7498
1.4363 | -90.6171
1.44 | -90.4508
1.4437 | -90.37
1.4475 | -90.4691
1.4512 | -90.6583
1.4549 | -90.7366
1.4586 | -90.6349
1.4623 | -90.489
1.466 | -90.4703
1.4697 | -90.6486
1.4734 | -91.025
1.4772 | -91.455
1.4809 | -91.5754
1.4846 | -91.3411
1.4883 | -91.1111
1.492 | -91.0276
1.4957 | -90.9902
1.4994 | -90.9215
1.5031 | -90.8484
1.5068 | -90.7612
1.5106 | -90.685
1.5143 | -90.7051
1.518 | -90.7782
1.5217 | -90.9222
1.5254 | -91.2562
1.5291 | -91.6607
1.5328 | -91.7994
1.5365 | -91.6492
1.5402 | -91.5658
1.544 | -91.7331
1.5477 | -92.0371
1.5514 | -92.2428
1.5551 | -92.0783
1.5588 | -91.5792
1.5625 | -91.0546
1.5662 | -90.7994
1.5699 | -90.9144
1.5736 | -91.2524
1.5774 | -91.5762
1.5811 | -91.8081
1.5848 | -91.9796
1.5885 | -91.9473
1.5922 | -91.5811
1.5959 | -91.2419
1.5996 | -91.3311
1.6033 | -91.767
1.6071 | -92.0422
1.6108 | -92.2009
1.6145 | -92.2598
1.6182 | -92.3085
1.6219 | -92.3816
1.6256 | -92.4974
1.6293 | -92.6503
1.633 | -92.748
1.6367 | -92.6427
1.6405 | -92.3373
1.6442 | -92.0313
1.6479 | -91.8693
1.6516 | -91.7917
1.6553 | -91.6773
1.659 | -91.53
1.6627 | -91.453
1.6664 | -91.4666
1.6701 | -91.5208
1.6739 | -91.4961
1.6776 | -91.358
1.6813 | -91.2275
1.685 | -91.256
1.6887 | -91.452
1.6924 | -91.6876
1.6961 | -91.868
1.6998 | -92.0227
1.7035 | -92.1507
1.7073 | -92.2167
1.711 | -92.2087
1.7147 | -92.1564
1.7184 | -92.1056
1.7221 | -92.0511
1.7258 | -92.0005
1.7295 | -92.0118
1.7332 | -92.1456
1.737 | -92.3996
1.7407 | -92.6795
1.7444 | -92.8552
1.7481 | -92.8926
1.7518 | -92.8212
1.7555 | -92.7454
1.7592 | -92.7284
1.7629 | -92.7488
1.7666 | -92.768
1.7704 | -92.7956
1.7741 | -92.8217
1.7778 | -92.7526
1.7815 | -92.5329
1.7852 | -92.2989
1.7889 | -92.2497
1.7926 | -92.4837
1.7963 | -92.911
1.8 | -93.3046
1.8038 | -93.4834
1.8075 | -93.4503
1.8112 | -93.3229
1.8149 | -93.1686
1.8186 | -93.001
1.8223 | -92.8231
1.826 | -92.622
1.8297 | -92.4474
1.8334 | -92.3881
1.8372 | -92.4899
1.8409 | -92.6665
1.8446 | -92.7568
1.8483 | -92.7273
1.852 | -92.6837
1.8557 | -92.7072
1.8594 | -92.8048
1.8631 | -92.9649
1.8669 | -93.1967
1.8706 | -93.4621
1.8743 | -93.6598
1.878 | -93.6843
1.8817 | -93.506
1.8854 | -93.2115
1.8891 | -92.9403
1.8928 | -92.782
1.8965 | -92.7651
1.9003 | -92.8559
1.904 | -92.983
1.9077 | -93.0675
1.9114 | -93.0527
1.9151 | -92.9428
1.9188 | -92.8036
1.9225 | -92.7034
1.9262 | -92.6969
1.9299 | -92.8002
1.9337 | -92.9928
1.9374 | -93.2285
1.9411 | -93.4692
1.9448 | -93.6893
1.9485 | -93.8295
1.9522 | -93.8135
1.9559 | -93.6697
1.9596 | -93.5053
1.9633 | -93.4219
1.9671 | -93.4217
1.9708 | -93.4186
1.9745 | -93.3336
1.9782 | -93.1898
1.9819 | -93.0692
1.9856 | -93.0027
1.9893 | -92.9546
1.993 | -92.9277
1.9968 | -92.9435
2.0005 | -93.0023
2.0042 | -93.0882
2.0079 | -93.2274
2.0116 | -93.4536
2.0153 | -93.6996
2.019 | -93.8369
2.0227 | -93.7781
2.0264 | -93.5381
2.0302 | -93.2241
2.0339 | -92.9853
2.0376 | -92.9273
2.0413 | -93.0279
2.045 | -93.154
2.0487 | -93.1966
2.0524 | -93.1676
2.0561 | -93.1553
2.0598 | -93.2006
2.0636 | -93.2899
2.0673 | -93.4166
2.071 | -93.5987
2.0747 | -93.8323
2.0784 | -94.0542
2.0821 | -94.1707
2.0858 | -94.1414
2.0895 | -94.0054
2.0932 | -93.8725
2.097 | -93.805
2.1007 | -93.7817
2.1044 | -93.7213
2.1081 | -93.5858
2.1118 | -93.4643
2.1155 | -93.4837
2.1192 | -93.6578
2.1229 | -93.8833
2.1267 | -94.0173
2.1304 | -94.0442
2.1341 | -94.0376
2.1378 | -94.0291
2.1415 | -93.8894
2.1452 | -93.7484
2.1489 | -93.6135
2.1526 | -93.547
2.1563 | -93.5857
2.1601 | -93.7128
2.1638 | -93.8748
2.1675 | -93.9976
2.1712 | -94.057
2.1749 | -94.0954
2.1786 | -94.1532
2.1823 | -94.2281
2.186 | -94.2756
2.1897 | -94.2183
2.1935 | -94.0538
2.1972 | -93.8477
2.2009 | -93.6835
2.2046 | -93.6175
2.2083 | -93.6479
2.212 | -93.7558
2.2157 | -93.9171
2.2194 | -94.1106
2.2231 | -94.2969
2.2269 | -94.4196
2.2306 | -94.4872
2.2343 | -94.5201
2.238 | -94.5474
2.2417 | -94.5956
2.2454 | -94.6724
2.2491 | -94.7698
2.2528 | -94.8292
2.2566 | -94.8215
2.2603 | -94.7426
2.264 | -94.6231
2.2677 | -94.5093
2.2714 | -94.4123
2.2751 | -94.3444
2.2788 | -94.3038
2.2825 | -94.2686
2.2862 | -94.2152
2.29 | -94.1571
2.2937 | -94.1345
2.2974 | -94.1932
2.3011 | -94.3545
2.3048 | -94.5719
2.3085 | -94.7578
2.3122 | -94.8531
2.3159 | -94.8508
2.3196 | -94.7811
2.3234 | -94.6918
2.3271 | -94.6174
2.3308 | -94.5655
2.3345 | -94.5511
2.3382 | -94.5817
2.3419 | -94.6458
2.3456 | -94.7349
2.3493 | -94.8432
2.3531 | -94.9667
2.3568 | -95.1108
2.3605 | -95.2594
2.3642 | -95.3831
2.3679 | -95.4439
2.3716 | -95.4232
2.3753 | -95.3069
2.379 | -95.0962
2.3827 | -94.8368
2.3865 | -94.579
2.3902 | -94.3931
2.3939 | -94.3337
2.3976 | -94.4304
2.4013 | -94.6775
2.405 | -95.0081
2.4087 | -95.3301
2.4124 | -95.5373
2.4161 | -95.5885
2.4199 | -95.5275
2.4236 | -95.4145
2.4273 | -95.3248
2.431 | -95.2872
2.4347 | -95.2999
2.4384 | -95.3174
2.4421 | -95.3131
2.4458 | -95.2657
2.4495 | -95.1767
2.4533 | -95.0714
2.457 | -94.9722
2.4607 | -94.9205
2.4644 | -94.9555
2.4681 | -95.0978
2.4718 | -95.3431
2.4755 | -95.6282
2.4792 | -95.8567
2.483 | -95.9571
2.4867 | -95.9414
2.4904 | -95.8318
2.4941 | -95.6418
2.4978 | -95.4135
2.5015 | -95.2216
2.5052 | -95.1142
2.5089 | -95.0915
2.5126 | -95.128
2.5164 | -95.2055
2.5201 | -95.3325
2.5238 | -95.489
2.5275 | -95.6343
2.5312 | -95.7077
2.5349 | -95.6722
2.5386 | -95.5594
2.5423 | -95.4415
2.546 | -95.3809
2.5498 | -95.3602
2.5535 | -95.3384
2.5572 | -95.2885
2.5609 | -95.2216
2.5646 | -95.173
2.5683 | -95.1788
2.572 | -95.2667
2.5757 | -95.428
2.5794 | -95.5872
2.5832 | -95.669
2.5869 | -95.6702
2.5906 | -95.647
2.5943 | -95.6442
2.598 | -95.7043
2.6017 | -95.8254
2.6054 | -95.9713
2.6091 | -96.0566
2.6129 | -96.0142
2.6166 | -95.8651
2.6203 | -95.7059
2.624 | -95.5751
2.6277 | -95.4458
2.6314 | -95.2962
2.6351 | -95.1503
2.6388 | -95.0689
2.6425 | -95.1158
2.6463 | -95.3057
2.65 | -95.5792
2.6537 | -95.8301
2.6574 | -95.9654
2.6611 | -95.9406
2.6648 | -95.8075
2.6685 | -95.6622
2.6722 | -95.5844
2.6759 | -95.6581
2.6797 | -95.7017
2.6834 | -95.7685
2.6871 | -95.8407
2.6908 | -95.9105
2.6945 | -95.9687
2.6982 | -96.0159
2.7019 | -96.0264
2.7056 | -96.0005
2.7093 | -95.9598
2.7131 | -95.911
2.7168 | -95.8617
2.7205 | -95.8103
2.7242 | -95.7625
2.7279 | -95.7162
2.7316 | -95.6699
2.7353 | -95.6246
2.739 | -95.581
2.7428 | -95.5488
2.7465 | -95.5169
2.7502 | -95.4876
2.7539 | -95.4523
2.7576 | -95.4163
2.7613 | -95.4243
2.765 | -95.4984
2.7687 | -95.6387
2.7724 | -95.8186
2.7762 | -96.0121
2.7799 | -96.1942
2.7836 | -96.3474
2.7873 | -96.4418
2.791 | -96.4522
2.7947 | -96.368
2.7984 | -96.2131
2.8021 | -96.0535
2.8058 | -95.9303
2.8096 | -95.871
2.8133 | -95.8932
2.817 | -95.9605
2.8207 | -96.0267
2.8244 | -96.0372
2.8281 | -95.9574
2.8318 | -95.8131
2.8355 | -95.664
2.8392 | -95.5506
2.843 | -95.5004
2.8467 | -95.5056
2.8504 | -95.5539
2.8541 | -95.6385
2.8578 | -95.7426
2.8615 | -95.8506
2.8652 | -95.9431
2.8689 | -95.9839
2.8727 | -95.9632
2.8764 | -95.9128
2.8801 | -95.8643
2.8838 | -95.8473
2.8875 | -95.8898
2.8912 | -95.9842
2.8949 | -96.1275
2.8986 | -96.284
2.9023 | -96.411
2.9061 | -96.4669
2.9098 | -96.4369
2.9135 | -96.3591
2.9172 | -96.2931
2.9209 | -96.2754
2.9246 | -96.2917
2.9283 | -96.3188
2.932 | -96.3317
2.9357 | -96.3047
2.9395 | -96.2597
2.9432 | -96.2047
2.9469 | -96.1742
2.9506 | -96.1988
2.9543 | -96.287
2.958 | -96.4253
2.9617 | -96.5915
2.9654 | -96.7229
2.9691 | -96.7764
2.9729 | -96.7505
2.9766 | -96.6556
2.9803 | -96.5414
2.984 | -96.4656
2.9877 | -96.4627
2.9914 | -96.5551
2.9951 | -96.7243
2.9988 | -96.9439
3.0026 | -97.1757
3.0063 | -97.3646
3.01 | -97.44
3.0137 | -97.4093
3.0174 | -97.2841
3.0211 | -97.0867
3.0248 | -96.8755
3.0285 | -96.6881
3.0322 | -96.5654
3.036 | -96.5277
3.0397 | -96.5927
3.0434 | -96.7494
3.0471 | -96.9346
3.0508 | -97.0771
3.0545 | -97.1222
3.0582 | -97.0701
3.0619 | -96.9451
3.0656 | -96.7982
3.0694 | -96.6697
3.0731 | -96.5851
3.0768 | -96.5604
3.0805 | -96.5763
3.0842 | -96.6221
3.0879 | -96.6356
3.0916 | -96.5888
3.0953 | -96.5163
3.099 | -96.4468
3.1028 | -96.4213
3.1065 | -96.4553
3.1102 | -96.5538
3.1139 | -96.7058
3.1176 | -96.8885
3.1213 | -97.0691
3.125 | -97.2197
3.1287 | -97.3342
3.1325 | -97.3851
3.1362 | -97.379
3.1399 | -97.3013
3.1436 | -97.1632
3.1473 | -97.0127
3.151 | -96.8698
3.1547 | -96.7613
3.1584 | -96.6817
3.1621 | -96.6401
3.1659 | -96.6372
3.1696 | -96.6696
3.1733 | -96.7308
3.177 | -96.8199
3.1807 | -96.9428
3.1844 | -97.1091
3.1881 | -97.3153
3.1918 | -97.5378
3.1955 | -97.7499
3.1993 | -97.9121
3.203 | -97.9728
3.2067 | -97.9427
3.2104 | -97.7884
3.2141 | -97.6772
3.2178 | -97.5481
3.2215 | -97.4235
3.2252 | -97.3159
3.2289 | -97.2331
3.2327 | -97.1871
3.2364 | -97.1684
3.2401 | -97.1681
3.2438 | -97.1712
3.2475 | -97.1627
3.2512 | -97.1522
3.2549 | -97.1594
3.2586 | -97.1893
3.2624 | -97.2522
3.2661 | -97.3408
3.2698 | -97.4433
3.2735 | -97.5247
3.2772 | -97.5492
3.2809 | -97.5028
3.2846 | -97.403
3.2883 | -97.2738
3.292 | -97.1573
3.2958 | -97.0736
3.2995 | -97.0313
3.3032 | -97.033
3.3069 | -97.0794
3.3106 | -97.1546
3.3143 | -97.2561
3.318 | -97.3592
3.3217 | -97.4552
3.3254 | -97.5286
3.3292 | -97.5903
3.3329 | -97.643
3.3366 | -97.6848
3.3403 | -97.7042
3.344 | -97.7008
3.3477 | -97.6854
3.3514 | -97.6775
3.3551 | -97.6915
3.3588 | -97.7328
3.3626 | -97.7903
3.3663 | -97.859
3.37 | -97.9163
3.3737 | -97.963
3.3774 | -97.9881
3.3811 | -97.9874
3.3848 | -97.9575
3.3885 | -97.9086
3.3923 | -97.8518
3.396 | -97.7934
3.3997 | -97.7556
3.4034 | -97.738
3.4071 | -97.743
3.4108 | -97.7707
3.4145 | -97.809
3.4182 | -97.8377
3.4219 | -97.8376
3.4257 | -97.8052
3.4294 | -97.7486
3.4331 | -97.686
3.4368 | -97.6415
3.4405 | -97.6226
3.4442 | -97.6116
3.4479 | -97.6044
3.4516 | -97.6073
3.4553 | -97.62
3.4591 | -97.661
3.4628 | -97.7436
3.4665 | -97.8747
3.4702 | -98.0431
3.4739 | -98.2181
3.4776 | -98.35
3.4813 | -98.4025
3.485 | -98.3581
3.4888 | -98.2261
3.4925 | -98.0357
3.4962 | -97.844
3.4999 | -97.7055
3.5036 | -97.6269
3.5073 | -97.5944
3.511 | -97.5952
3.5147 | -97.5933
3.5184 | -97.5848
3.5222 | -97.5527
3.5259 | -97.5098
3.5296 | -97.4914
3.5333 | -97.5099
3.537 | -97.5678
3.5407 | -97.6707
3.5444 | -97.8089
3.5481 | -97.9685
3.5518 | -98.1206
3.5556 | -98.2482
3.5593 | -98.3482
3.563 | -98.4187
3.5667 | -98.4441
3.5704 | -98.4056
3.5741 | -98.3036
3.5778 | -98.1799
3.5815 | -98.0287
3.5852 | -97.8631
3.589 | -97.7284
3.5927 | -97.6458
3.5964 | -97.6282
3.6001 | -97.6779
3.6038 | -97.7813
3.6075 | -97.9263
3.6112 | -98.0761
3.6149 | -98.1927
3.6187 | -98.2461
3.6224 | -98.2353
3.6261 | -98.1912
3.6298 | -98.1332
3.6335 | -98.0932
3.6372 | -98.1253
3.6409 | -98.2192
3.6446 | -98.3544
3.6483 | -98.5037
3.6521 | -98.6262
3.6558 | -98.7036
3.6595 | -98.7328
3.6632 | -98.7066
3.6669 | -98.6606
3.6706 | -98.6094
3.6743 | -98.5608
3.678 | -98.5119
3.6817 | -98.4681
3.6855 | -98.4459
3.6892 | -98.455
3.6929 | -98.4888
3.6966 | -98.5406
3.7003 | -98.5765
3.704 | -98.5724
3.7077 | -98.5196
3.7114 | -98.4338
3.7151 | -98.3371
3.7189 | -98.2698
3.7226 | -98.2368
3.7263 | -98.2322
3.73 | -98.2468
3.7337 | -98.2804
3.7374 | -98.3225
3.7411 | -98.3583
3.7448 | -98.3734
3.7485 | -98.3953
3.7523 | -98.4094
3.756 | -98.421
3.7597 | -98.4331
3.7634 | -98.4469
3.7671 | -98.4758
3.7708 | -98.5125
3.7745 | -98.5624
3.7782 | -98.6159
3.782 | -98.6847
3.7857 | -98.7748
3.7894 | -98.8654
3.7931 | -98.9405
3.7968 | -98.9987
3.8005 | -99.0384
3.8042 | -99.0574
3.8079 | -99.0566
3.8116 | -99.0394
3.8154 | -99.0159
3.8191 | -98.9863
3.8228 | -98.9691
3.8265 | -98.9703
3.8302 | -98.9929
3.8339 | -99.0505
3.8376 | -99.142
3.8413 | -99.2609
3.845 | -99.3815
3.8488 | -99.4889
3.8525 | -99.556
3.8562 | -99.5702
3.8599 | -99.5294
3.8636 | -99.444
3.8673 | -99.3324
3.871 | -99.2119
3.8747 | -99.1059
3.8784 | -99.0319
3.8822 | -98.9971
3.8859 | -99.0029
3.8896 | -99.052
3.8933 | -99.1307
3.897 | -99.2167
3.9007 | -99.3081
3.9044 | -99.4019
3.9081 | -99.4777
3.9119 | -99.5153
3.9156 | -99.53
3.9193 | -99.509
3.923 | -99.4472
3.9267 | -99.3604
3.9304 | -99.2613
3.9341 | -99.1794
3.9378 | -99.1267
3.9415 | -99.1068
3.9453 | -99.1219
3.949 | -99.1701
3.9527 | -99.2271
3.9564 | -99.2787
3.9601 | -99.3097
3.9638 | -99.3007
3.9675 | -99.2471
3.9712 | -99.1681
3.9749 | -99.0732
3.9787 | -98.9744
3.9824 | -98.8868
3.9861 | -98.8218
3.9898 | -98.7794
3.9935 | -98.7639
3.9972 | -98.761
4.0009 | -98.7737
4.0046 | -98.805
4.0084 | -98.8477
4.0121 | -98.891
4.0158 | -98.9228
4.0195 | -98.9411
4.0232 | -98.9401
4.0269 | -98.9211
4.0306 | -98.8937
4.0343 | -98.8615
4.038 | -98.832
4.0418 | -98.7977
4.0455 | -98.7635
4.0492 | -98.7476
4.0529 | -98.747
4.0566 | -98.7612
4.0603 | -98.7913
4.064 | -98.8311
4.0677 | -98.8716
4.0714 | -98.9078
4.0752 | -98.9308
4.0789 | -98.9444
4.0826 | -98.9493
4.0863 | -98.9455
4.09 | -98.9353
4.0937 | -98.9212
4.0974 | -98.9114
4.1011 | -98.9031
4.1048 | -98.9042
4.1086 | -98.9243
4.1123 | -98.9589
4.116 | -99.0138
4.1197 | -99.0843
4.1234 | -99.1626
4.1271 | -99.2414
4.1308 | -99.317
4.1345 | -99.3836
4.1383 | -99.4368
4.142 | -99.4698
4.1457 | -99.4664
4.1494 | -99.4426
4.1531 | -99.3975
4.1568 | -99.338
4.1605 | -99.272
4.1642 | -99.2108
4.1679 | -99.1682
4.1717 | -99.1401
4.1754 | -99.1333
4.1791 | -99.1571
4.1828 | -99.2044
4.1865 | -99.2735
4.1902 | -99.3533
4.1939 | -99.4352
4.1976 | -99.5156
4.2013 | -99.5889
4.2051 | -99.6477
4.2088 | -99.6889
4.2125 | -99.7014
4.2162 | -99.7002
4.2199 | -99.6923
4.2236 | -99.6866
4.2273 | -99.6925
4.231 | -99.6991
4.2347 | -99.7091
4.2385 | -99.716
4.2422 | -99.7149
4.2459 | -99.7175
4.2496 | -99.7112
4.2533 | -99.6931
4.257 | -99.6715
4.2607 | -99.6439
4.2644 | -99.6133
4.2682 | -99.5921
4.2719 | -99.5856
4.2756 | -99.6007
4.2793 | -99.6317
4.283 | -99.6729
4.2867 | -99.7021
4.2904 | -99.7224
4.2941 | -99.7316
4.2978 | -99.7271
4.3016 | -99.7177
4.3053 | -99.7079
4.309 | -99.7036
4.3127 | -99.7056
4.3164 | -99.7243
4.3201 | -99.7549
4.3238 | -99.785
4.3275 | -99.8177
4.3312 | -99.8584
4.335 | -99.8988
4.3387 | -99.934
4.3424 | -99.9587
4.3461 | -99.9872
4.3498 | -100.0085
4.3535 | -100.0265
4.3572 | -100.0512
4.3609 | -100.0819
4.3646 | -100.0966
4.3684 | -100.1045
4.3721 | -100.1002
4.3758 | -100.0712
4.3795 | -100.0288
4.3832 | -99.9775
4.3869 | -99.9175
4.3906 | -99.8579
4.3943 | -99.7988
4.3981 | -99.7369
4.4018 | -99.6777
4.4055 | -99.635
4.4092 | -99.6127
4.4129 | -99.6068
4.4166 | -99.6094
4.4203 | -99.6355
4.424 | -99.6777
4.4277 | -99.7194
4.4315 | -99.7669
4.4352 | -99.8247
4.4389 | -99.8735
4.4426 | -99.9112
4.4463 | -99.9406
4.45 | -99.9473
4.4537 | -99.9424
4.4574 | -99.9351
4.4611 | -99.9241
4.4649 | -99.922
4.4686 | -99.9356
4.4723 | -99.9629
4.476 | -99.9988
4.4797 | -100.0411
4.4834 | -100.0895
4.4871 | -100.1359
4.4908 | -100.1773
4.4945 | -100.2189
4.4983 | -100.2547
4.502 | -100.2766
4.5057 | -100.2858
4.5094 | -100.2779
4.5131 | -100.2559
4.5168 | -100.2216
4.5205 | -100.1901
4.5242 | -100.1705
4.528 | -100.1701
4.5317 | -100.204
4.5354 | -100.2609
4.5391 | -100.3369
4.5428 | -100.424
4.5465 | -100.5075
4.5502 | -100.5805
4.5539 | -100.6317
4.5576 | -100.6524
4.5614 | -100.6399
4.5651 | -100.5911
4.5688 | -100.522
4.5725 | -100.4316
4.5762 | -100.3304
4.5799 | -100.2304
4.5836 | -100.1433
4.5873 | -100.0823
4.591 | -100.0432
4.5948 | -100.03
4.5985 | -100.0433
4.6022 | -100.071
4.6059 | -100.1132
4.6096 | -100.1601
4.6133 | -100.2103
4.617 | -100.2631
4.6207 | -100.3102
4.6244 | -100.3608
4.6282 | -100.4146
4.6319 | -100.472
4.6356 | -100.5184
4.6393 | -100.5494
4.643 | -100.5587
4.6467 | -100.547
4.6504 | -100.5147
4.6541 | -100.4604
4.6579 | -100.3994
4.6616 | -100.3488
4.6653 | -100.3065
4.669 | -100.2802
4.6727 | -100.2787
4.6764 | -100.2888
4.6801 | -100.306
4.6838 | -100.3286
4.6875 | -100.3513
4.6913 | -100.3766
4.695 | -100.3904
4.6987 | -100.4002
4.7024 | -100.3991
4.7061 | -100.3981
4.7098 | -100.4083
4.7135 | -100.413
4.7172 | -100.407
4.7209 | -100.3997
4.7247 | -100.3913
4.7284 | -100.378
4.7321 | -100.3495
4.7358 | -100.3232
4.7395 | -100.2911
4.7432 | -100.2613
4.7469 | -100.2447
4.7506 | -100.2298
4.7543 | -100.2117
4.7581 | -100.1915
4.7618 | -100.1685
4.7655 | -100.1556
4.7692 | -100.1472
4.7729 | -100.1467
4.7766 | -100.1601
4.7803 | -100.1688
4.784 | -100.1864
4.7878 | -100.2154
4.7915 | -100.243
4.7952 | -100.2663
4.7989 | -100.2891
4.8026 | -100.3182
4.8063 | -100.3402
4.81 | -100.3633
4.8137 | -100.3774
4.8174 | -100.3723
4.8212 | -100.3681
4.8249 | -100.3639
4.8286 | -100.3519
4.8323 | -100.3369
4.836 | -100.312
4.8397 | -100.2843
4.8434 | -100.257
4.8471 | -100.2347
4.8508 | -100.2176
4.8546 | -100.2159
4.8583 | -100.2237
4.862 | -100.2333
4.8657 | -100.2559
4.8694 | -100.2917
4.8731 | -100.327
4.8768 | -100.3626
4.8805 | -100.4028
4.8842 | -100.4423
4.888 | -100.4804
4.8917 | -100.5054
4.8954 | -100.5113
4.8991 | -100.5024
4.9028 | -100.4836
4.9065 | -100.4571
4.9102 | -100.4253
4.9139 | -100.399
4.9177 | -100.3785
4.9214 | -100.3697
4.9251 | -100.3697
4.9288 | -100.3788
4.9325 | -100.4002
4.9362 | -100.439
4.9399 | -100.485
4.9436 | -100.531
4.9473 | -100.5759
4.9511 | -100.6237
4.9548 | -100.664
4.9585 | -100.6925
4.9622 | -100.7075
4.9659 | -100.7191
4.9696 | -100.7334
4.9733 | -100.753
4.977 | -100.7669
4.9807 | -100.7803
4.9845 | -100.7942
4.9882 | -100.7993
4.9919 | -100.7922
4.9956 | -100.7721
4.9993 | -100.7487
5.003 | -100.7235
5.0067 | -100.6838
5.0104 | -100.6397
5.0141 | -100.5907
5.0179 | -100.5473
5.0216 | -100.5185
5.0253 | -100.5014
5.029 | -100.4965
5.0327 | -100.5016
5.0364 | -100.5196
5.0401 | -100.5479
5.0438 | -100.5763
5.0476 | -100.6148
5.0513 | -100.65
5.055 | -100.6904
5.0587 | -100.7255
5.0624 | -100.7538
5.0661 | -100.7867
5.0698 | -100.8147
5.0735 | -100.8397
5.0772 | -100.8679
5.081 | -100.9018
5.0847 | -100.9436
5.0884 | -100.9996
5.0921 | -101.0601
5.0958 | -101.1163
5.0995 | -101.1568
5.1032 | -101.18
5.1069 | -101.1821
5.1106 | -101.1667
5.1144 | -101.14
5.1181 | -101.1095
5.1218 | -101.0839
5.1255 | -101.0571
5.1292 | -101.0385
5.1329 | -101.0321
5.1366 | -101.0438
5.1403 | -101.0568
5.1441 | -101.0584
5.1478 | -101.0529
5.1515 | -101.0313
5.1552 | -100.9952
5.1589 | -100.9424
5.1626 | -100.8833
5.1663 | -100.8086
5.17 | -100.7312
5.1737 | -100.6624
5.1775 | -100.6093
5.1812 | -100.5796
5.1849 | -100.5736
5.1886 | -100.5942
5.1923 | -100.6322
5.196 | -100.6946
5.1997 | -100.774
5.2034 | -100.8754
5.2071 | -100.9845
5.2109 | -101.0904
5.2146 | -101.1805
5.2183 | -101.2545
5.222 | -101.3007
5.2257 | -101.318
5.2294 | -101.3107
5.2331 | -101.2741
5.2368 | -101.2258
5.2405 | -101.1647
5.2443 | -101.1094
5.248 | -101.0673
5.2517 | -101.0296
5.2554 | -101.0004
5.2591 | -100.9798
5.2628 | -100.9657
5.2665 | -100.957
5.2702 | -100.9549
5.274 | -100.9573
5.2777 | -100.9606
5.2814 | -100.9633
5.2851 | -100.9743
5.2888 | -100.9932
5.2925 | -101.0274
5.2962 | -101.0756
5.2999 | -101.1343
5.3036 | -101.1931
5.3074 | -101.2538
5.3111 | -101.3106
5.3148 | -101.3565
5.3185 | -101.3971
5.3222 | -101.4131
5.3259 | -101.4023
5.3296 | -101.3822
5.3333 | -101.3428
5.337 | -101.2857
5.3408 | -101.2257
5.3445 | -101.1702
5.3482 | -101.1251
5.3519 | -101.1218
5.3556 | -101.1063
5.3593 | -101.1086
5.363 | -101.1257
5.3667 | -101.1552
5.3704 | -101.1919
5.3742 | -101.2275
5.3779 | -101.2575
5.3816 | -101.288
5.3853 | -101.3187
5.389 | -101.3393
5.3927 | -101.3624
5.3964 | -101.3808
5.4001 | -101.3892
5.4039 | -101.4015
5.4076 | -101.4192
5.4113 | -101.4349
5.415 | -101.459
5.4187 | -101.4853
5.4224 | -101.5123
5.4261 | -101.5396
5.4298 | -101.5686
5.4335 | -101.5983
5.4373 | -101.6221
5.441 | -101.636
5.4447 | -101.649
5.4484 | -101.6556
5.4521 | -101.6494
5.4558 | -101.6305
5.4595 | -101.6074
5.4632 | -101.5716
5.4669 | -101.5264
5.4707 | -101.4755
5.4744 | -101.4288
5.4781 | -101.3853
5.4818 | -101.3527
5.4855 | -101.3355
5.4892 | -101.3238
5.4929 | -101.3253
5.4966 | -101.3411
5.5003 | -101.369
5.5041 | -101.4002
5.5078 | -101.4283
5.5115 | -101.4493
5.5152 | -101.4568
5.5189 | -101.4533
5.5226 | -101.4308
5.5263 | -101.3913
5.53 | -101.3405
5.5338 | -101.2893
5.5375 | -101.2545
5.5412 | -101.2321
5.5449 | -101.2179
5.5486 | -101.2256
5.5523 | -101.2478
5.556 | -101.2817
5.5597 | -101.3226
5.5634 | -101.3696
5.5672 | -101.4128
5.5709 | -101.4448
5.5746 | -101.4711
5.5783 | -101.4819
5.582 | -101.4772
5.5857 | -101.4614
5.5894 | -101.4432
5.5931 | -101.42
5.5968 | -101.4008
5.6006 | -101.3925
5.6043 | -101.4005
5.608 | -101.4255
5.6117 | -101.4717
5.6154 | -101.5381
5.6191 | -101.6183
5.6228 | -101.7033
5.6265 | -101.7981
5.6302 | -101.8869
5.634 | -101.9502
5.6377 | -101.9886
5.6414 | -102.0085
5.6451 | -102.0079
5.6488 | -101.9846
5.6525 | -101.9538
5.6562 | -101.9156
5.6599 | -101.8729
5.6637 | -101.8367
5.6674 | -101.8119
5.6711 | -101.8038
5.6748 | -101.8084
5.6785 | -101.8232
5.6822 | -101.8476
5.6859 | -101.8871
5.6896 | -101.9304
5.6933 | -101.9822
5.6971 | -102.0316
5.7008 | -102.0741
5.7045 | -102.107
5.7082 | -102.132
5.7119 | -102.1559
5.7156 | -102.1719
5.7193 | -102.178
5.723 | -102.1672
5.7267 | -102.1375
5.7305 | -102.0942
5.7342 | -102.0502
5.7379 | -102.0035
5.7416 | -101.963
5.7453 | -101.9325
5.749 | -101.914
5.7527 | -101.9104
5.7564 | -101.9176
5.7601 | -101.9372
5.7639 | -101.9595
5.7676 | -101.9762
5.7713 | -101.9793
5.775 | -101.9753
5.7787 | -101.9632
5.7824 | -101.9473
5.7861 | -101.931
5.7898 | -101.9203
5.7936 | -101.9048
5.7973 | -101.8801
5.801 | -101.8652
5.8047 | -101.8547
5.8084 | -101.8438
5.8121 | -101.8361
5.8158 | -101.8293
5.8195 | -101.8131
5.8232 | -101.7833
5.827 | -101.7455
5.8307 | -101.6954
5.8344 | -101.6299
5.8381 | -101.5535
5.8418 | -101.4752
5.8455 | -101.3959
5.8492 | -101.3258
5.8529 | -101.2697
5.8566 | -101.234
5.8604 | -101.2208
5.8641 | -101.2241
5.8678 | -101.2452
5.8715 | -101.2813
5.8752 | -101.3377
5.8789 | -101.4073
5.8826 | -101.4788
5.8863 | -101.5373
5.89 | -101.5945
5.8938 | -101.6405
5.8975 | -101.6862
5.9012 | -101.7202
5.9049 | -101.7398
5.9086 | -101.7531
5.9123 | -101.7633
5.916 | -101.7607
5.9197 | -101.7515
5.9235 | -101.7443
5.9272 | -101.7403
5.9309 | -101.7373
5.9346 | -101.7324
5.9383 | -101.7242
5.942 | -101.7233
5.9457 | -101.7286
5.9494 | -101.7392
5.9531 | -101.7538
5.9569 | -101.7747
5.9606 | -101.8011
5.9643 | -101.8273
5.968 | -101.8608
5.9717 | -101.8994
5.9754 | -101.9316
5.9791 | -101.9619
5.9828 | -101.9899
5.9865 | -102.0211
5.9903 | -102.0473
5.994 | -102.0781
5.9977 | -102.1234
6.0014 | -102.1782
6.0051 | -102.2405
6.0088 | -102.3087
6.0125 | -102.379
6.0162 | -102.4402
6.0199 | -102.4966
6.0237 | -102.5469
6.0274 | -102.5919
6.0311 | -102.6242
6.0348 | -102.6307
6.0385 | -102.6229
6.0422 | -102.6059
6.0459 | -102.5833
6.0496 | -102.5497
6.0534 | -102.5113
6.0571 | -102.473
6.0608 | -102.4281
6.0645 | -102.3848
6.0682 | -102.3505
6.0719 | -102.318
6.0756 | -102.2881
6.0793 | -102.2598
6.083 | -102.2395
6.0868 | -102.2305
6.0905 | -102.2272
6.0942 | -102.2326
6.0979 | -102.2486
6.1016 | -102.2734
6.1053 | -102.2939
6.109 | -102.3107
6.1127 | -102.3148
6.1164 | -102.3071
6.1202 | -102.2818
6.1239 | -102.2442
6.1276 | -102.1958
6.1313 | -102.1368
6.135 | -102.0744
6.1387 | -102.0188
6.1424 | -101.9644
6.1461 | -101.9115
6.1498 | -101.8693
6.1536 | -101.8389
6.1573 | -101.8204
6.161 | -101.8119
6.1647 | -101.8166
6.1684 | -101.8352
6.1721 | -101.8638
6.1758 | -101.8986
6.1795 | -101.94
6.1833 | -101.9891
6.187 | -102.0382
6.1907 | -102.0854
6.1944 | -102.1283
6.1981 | -102.1693
6.2018 | -102.2055
6.2055 | -102.2332
6.2092 | -102.2616
6.2129 | -102.2943
6.2167 | -102.3271
6.2204 | -102.3642
6.2241 | -102.4083
6.2278 | -102.4502
6.2315 | -102.4906
6.2352 | -102.529
6.2389 | -102.5666
6.2426 | -102.593
6.2463 | -102.6207
6.2501 | -102.6378
6.2538 | -102.6419
6.2575 | -102.6373
6.2612 | -102.6296
6.2649 | -102.6214
6.2686 | -102.6073
6.2723 | -102.59
6.276 | -102.5718
6.2797 | -102.56
6.2835 | -102.5437
6.2872 | -102.5353
6.2909 | -102.5314
6.2946 | -102.5328
6.2983 | -102.5374
6.302 | -102.5393
6.3057 | -102.5445
6.3094 | -102.5457
6.3132 | -102.5367
6.3169 | -102.5177
6.3206 | -102.4973
6.3243 | -102.4653
6.328 | -102.4332
6.3317 | -102.4064
6.3354 | -102.3801
6.3391 | -102.3551
6.3428 | -102.3472
6.3466 | -102.3477
6.3503 | -102.3555
6.354 | -102.3686
6.3577 | -102.3754
6.3614 | -102.3822
6.3651 | -102.3802
6.3688 | -102.3673
6.3725 | -102.3475
6.3762 | -102.3352
6.38 | -102.3142
6.3837 | -102.2972
6.3874 | -102.2902
6.3911 | -102.2875
6.3948 | -102.2947
6.3985 | -102.3062
6.4022 | -102.3258
6.4059 | -102.3548
6.4097 | -102.3814
6.4134 | -102.4044
6.4171 | -102.4258
6.4208 | -102.4376
6.4245 | -102.4071
6.4282 | -102.3962
6.4319 | -102.3827
6.4356 | -102.3637
6.4393 | -102.3349
6.4431 | -102.3064
6.4468 | -102.2788
6.4505 | -102.2573
6.4542 | -102.2509
6.4579 | -102.2503
6.4616 | -102.2624
6.4653 | -102.2834
6.469 | -102.3073
6.4727 | -102.3288
6.4765 | -102.3609
6.4802 | -102.3961
6.4839 | -102.4191
6.4876 | -102.4391
6.4913 | -102.4502
6.495 | -102.4603
6.4987 | -102.4722
6.5024 | -102.479
6.5061 | -102.4855
6.5099 | -102.4928
6.5136 | -102.5008
6.5173 | -102.5102
6.521 | -102.5251
6.5247 | -102.5368
6.5284 | -102.5489
6.5321 | -102.5622
6.5358 | -102.5799
6.5396 | -102.6015
6.5433 | -102.6238
6.547 | -102.6433
6.5507 | -102.6577
6.5544 | -102.673
6.5581 | -102.6934
6.5618 | -102.7128
6.5655 | -102.7319
6.5692 | -102.7572
6.573 | -102.7873
6.5767 | -102.813
6.5804 | -102.8368
6.5841 | -102.8726
6.5878 | -102.9113
6.5915 | -102.9477
6.5952 | -102.9722
6.5989 | -102.9934
6.6026 | -103.0143
6.6064 | -103.0267
6.6101 | -103.0313
6.6138 | -103.0221
6.6175 | -103.0057
6.6212 | -102.9845
6.6249 | -102.9585
6.6286 | -102.9208
6.6323 | -102.8845
6.636 | -102.8365
6.6398 | -102.779
6.6435 | -102.7237
6.6472 | -102.6699
6.6509 | -102.6187
6.6546 | -102.5668
6.6583 | -102.5166
6.662 | -102.4714
6.6657 | -102.4355
6.6695 | -102.405
6.6732 | -102.3839
6.6769 | -102.3763
6.6806 | -102.382
6.6843 | -102.3972
6.688 | -102.424
6.6917 | -102.4607
6.6954 | -102.5148
6.6991 | -102.5729
6.7029 | -102.6386
6.7066 | -102.7137
6.7103 | -102.7887
6.714 | -102.8773
6.7177 | -102.9579
6.7214 | -103.0298
6.7251 | -103.0999
6.7288 | -103.1613
6.7325 | -103.1988
6.7363 | -103.2197
6.74 | -103.2292
6.7437 | -103.2179
6.7474 | -103.1891
6.7511 | -103.1388
6.7548 | -103.0793
6.7585 | -103.0211
6.7622 | -102.9589
6.7659 | -102.8923
6.7697 | -102.8302
6.7734 | -102.7709
6.7771 | -102.7167
6.7808 | -102.6696
6.7845 | -102.6266
6.7882 | -102.5876
6.7919 | -102.5517
6.7956 | -102.5166
6.7993 | -102.4855
6.8031 | -102.4626
6.8068 | -102.4502
6.8105 | -102.4493
6.8142 | -102.4498
6.8179 | -102.463
6.8216 | -102.488
6.8253 | -102.5227
6.829 | -102.5727
6.8328 | -102.6357
6.8365 | -102.7067
6.8402 | -102.7782
6.8439 | -102.8508
6.8476 | -102.9222
6.8513 | -102.9963
6.855 | -103.066
6.8587 | -103.1195
6.8624 | -103.1619
6.8662 | -103.1875
6.8699 | -103.1955
6.8736 | -103.1913
6.8773 | -103.1696
6.881 | -103.1381
6.8847 | -103.098
6.8884 | -103.0554
6.8921 | -103.0186
6.8958 | -102.9797
6.8996 | -102.9328
6.9033 | -102.8815
6.907 | -102.826
6.9107 | -102.77
6.9144 | -102.717
6.9181 | -102.6619
6.9218 | -102.6092
6.9255 | -102.5608
6.9293 | -102.5129
6.933 | -102.4787
6.9367 | -102.4552
6.9404 | -102.4389
6.9441 | -102.4334
6.9478 | -102.4395
6.9515 | -102.4515
6.9552 | -102.4679
6.9589 | -102.4966
6.9627 | -102.5133
6.9664 | -102.5327
6.9701 | -102.5524
6.9738 | -102.5703
6.9775 | -102.5843
6.9812 | -102.6021
6.9849 | -102.6169
6.9886 | -102.6321
6.9923 | -102.6476
6.9961 | -102.6639
6.9998 | -102.6819
7.0035 | -102.7067
7.0072 | -102.7332
7.0109 | -102.7644
7.0146 | -102.7944
7.0183 | -102.8242
7.022 | -102.8525
7.0257 | -102.8791
7.0295 | -102.9006
7.0332 | -102.9188
7.0369 | -102.9272
7.0406 | -102.9308
7.0443 | -102.9288
7.048 | -102.9203
7.0517 | -102.9047
7.0554 | -102.8875
7.0592 | -102.8657
7.0629 | -102.8441
7.0666 | -102.8267
7.0703 | -102.8191
7.074 | -102.8179
7.0777 | -102.8246
7.0814 | -102.8384
7.0851 | -102.8589
7.0888 | -102.888
7.0926 | -102.9281
7.0963 | -102.9753
7.1 | -103.0269
7.1037 | -103.0792
7.1074 | -103.1414
7.1111 | -103.2012
7.1148 | -103.2541
7.1185 | -103.3006
7.1222 | -103.3475
7.126 | -103.3909
7.1297 | -103.4288
7.1334 | -103.4562
7.1371 | -103.4792
7.1408 | -103.4962
7.1445 | -103.5081
7.1482 | -103.5096
7.1519 | -103.5133
7.1556 | -103.5125
7.1594 | -103.5095
7.1631 | -103.5021
7.1668 | -103.5028
7.1705 | -103.5057
7.1742 | -103.5012
7.1779 | -103.4971
7.1816 | -103.49
7.1853 | -103.4814
7.1891 | -103.4651
7.1928 | -103.438
7.1965 | -103.4003
7.2002 | -103.3566
7.2039 | -103.3102
7.2076 | -103.2606
7.2113 | -103.2074
7.215 | -103.1518
7.2187 | -103.1062
7.2225 | -103.0635
7.2262 | -103.0223
7.2299 | -102.9866
7.2336 | -102.9569
7.2373 | -102.937
7.241 | -102.9242
7.2447 | -102.9186
7.2484 | -102.9248
7.2521 | -102.9367
7.2559 | -102.9618
7.2596 | -102.9854
7.2633 | -103.0141
7.267 | -103.0482
7.2707 | -103.0856
7.2744 | -103.1186
7.2781 | -103.1516
7.2818 | -103.1831
7.2855 | -103.2096
7.2893 | -103.2352
7.293 | -103.2545
7.2967 | -103.2677
7.3004 | -103.2804
7.3041 | -103.2923
7.3078 | -103.3003
7.3115 | -103.3044
7.3152 | -103.3053
7.319 | -103.3039
7.3227 | -103.3024
7.3264 | -103.2969
7.3301 | -103.292
7.3338 | -103.2833
7.3375 | -103.27
7.3412 | -103.2592
7.3449 | -103.2521
7.3486 | -103.2502
7.3524 | -103.2477
7.3561 | -103.2522
7.3598 | -103.2678
7.3635 | -103.2853
7.3672 | -103.3062
7.3709 | -103.3264
7.3746 | -103.3476
7.3783 | -103.3637
7.382 | -103.3769
7.3858 | -103.3876
7.3895 | -103.3957
7.3932 | -103.3993
7.3969 | -103.398
7.4006 | -103.4014
7.4043 | -103.408
7.408 | -103.4101
7.4117 | -103.4016
7.4154 | -103.3914
7.4192 | -103.3884
7.4229 | -103.3855
7.4266 | -103.3838
7.4303 | -103.3845
7.434 | -103.3908
7.4377 | -103.4048
7.4414 | -103.4269
7.4451 | -103.4516
7.4489 | -103.4764
7.4526 | -103.5041
7.4563 | -103.5273
7.46 | -103.5465
7.4637 | -103.5597
7.4674 | -103.5702
7.4711 | -103.5723
7.4748 | -103.5626
7.4785 | -103.5407
7.4823 | -103.5054
7.486 | -103.4625
7.4897 | -103.4103
7.4934 | -103.3562
7.4971 | -103.3042
7.5008 | -103.26
7.5045 | -103.2249
7.5082 | -103.1964
7.5119 | -103.1713
7.5157 | -103.1562
7.5194 | -103.1516
7.5231 | -103.1584
7.5268 | -103.1696
7.5305 | -103.1873
7.5342 | -103.2138
7.5379 | -103.2434
7.5416 | -103.2763
7.5453 | -103.313
7.5491 | -103.3517
7.5528 | -103.3888
7.5565 | -103.4259
7.5602 | -103.4644
7.5639 | -103.5003
7.5676 | -103.5364
7.5713 | -103.57
7.575 | -103.5928
7.5788 | -103.6073
7.5825 | -103.6228
7.5862 | -103.6343
7.5899 | -103.6404
7.5936 | -103.6402
7.5973 | -103.6358
7.601 | -103.6238
7.6047 | -103.6072
7.6084 | -103.5898
7.6122 | -103.5817
7.6159 | -103.5693
7.6196 | -103.5574
7.6233 | -103.5551
7.627 | -103.559
7.6307 | -103.5651
7.6344 | -103.579
7.6381 | -103.5957
7.6418 | -103.6169
7.6456 | -103.6322
7.6493 | -103.6478
7.653 | -103.6648
7.6567 | -103.6762
7.6604 | -103.6821
7.6641 | -103.6889
7.6678 | -103.7001
7.6715 | -103.7091
7.6752 | -103.7099
7.679 | -103.7072
7.6827 | -103.7021
7.6864 | -103.6943
7.6901 | -103.6832
7.6938 | -103.6659
7.6975 | -103.6414
7.7012 | -103.6157
7.7049 | -103.5859
7.7087 | -103.5503
7.7124 | -103.5131
7.7161 | -103.4765
7.7198 | -103.4389
7.7235 | -103.3996
7.7272 | -103.3604
7.7309 | -103.3243
7.7346 | -103.291
7.7383 | -103.2665
7.7421 | -103.2485
7.7458 | -103.2364
7.7495 | -103.2284
7.7532 | -103.2271
7.7569 | -103.2336
7.7606 | -103.2472
7.7643 | -103.2652
7.768 | -103.2904
7.7717 | -103.3259
7.7755 | -103.3635
7.7792 | -103.4015
7.7829 | -103.4426
7.7866 | -103.4837
7.7903 | -103.5328
7.794 | -103.5823
7.7977 | -103.6262
7.8014 | -103.6652
7.8052 | -103.6986
7.8089 | -103.7296
7.8126 | -103.7589
7.8163 | -103.7803
7.82 | -103.79
7.8237 | -103.7962
7.8274 | -103.7979
7.8311 | -103.7988
7.8348 | -103.7943
7.8386 | -103.7876
7.8423 | -103.7791
7.846 | -103.7697
7.8497 | -103.7603
7.8534 | -103.7397
7.8571 | -103.7205
7.8608 | -103.6988
7.8645 | -103.672
7.8682 | -103.6403
7.872 | -103.6051
7.8757 | -103.5728
7.8794 | -103.5397
7.8831 | -103.5082
7.8868 | -103.4766
7.8905 | -103.4533
7.8942 | -103.432
7.8979 | -103.4164
7.9016 | -103.4111
7.9054 | -103.4104
7.9091 | -103.4148
7.9128 | -103.4316
7.9165 | -103.4585
7.9202 | -103.4909
7.9239 | -103.5306
7.9276 | -103.5739
7.9313 | -103.6205
7.9351 | -103.6708
7.9388 | -103.7206
7.9425 | -103.7673
7.9462 | -103.8077
7.9499 | -103.8474
7.9536 | -103.88
7.9573 | -103.9081
7.961 | -103.9342
7.9647 | -103.9556
7.9685 | -103.9778
7.9722 | -103.9935
7.9759 | -104.0107
7.9796 | -104.0302
7.9833 | -104.0545
7.987 | -104.0749
7.9907 | -104.0884
7.9944 | -104.1041
7.9981 | -104.1195
8.0019 | -104.1266
8.0056 | -104.1271
8.0093 | -104.1271
8.013 | -104.1126
8.0167 | -104.0893
8.0204 | -104.0578
8.0241 | -104.0214
8.0278 | -103.9791
8.0315 | -103.9387
8.0353 | -103.8995
8.039 | -103.8666
8.0427 | -103.8377
8.0464 | -103.8198
8.0501 | -103.8136
8.0538 | -103.8207
8.0575 | -103.8342
8.0612 | -103.858
8.0649 | -103.8942
8.0687 | -103.9375
8.0724 | -103.985
8.0761 | -104.0321
8.0798 | -104.081
8.0835 | -104.1298
8.0872 | -104.1777
8.0909 | -104.221
8.0946 | -104.2542
8.0984 | -104.2735
8.1021 | -104.2842
8.1058 | -104.2896
8.1095 | -104.2889
8.1132 | -104.2768
8.1169 | -104.2566
8.1206 | -104.2391
8.1243 | -104.2158
8.128 | -104.1822
8.1318 | -104.1538
8.1355 | -104.1267
8.1392 | -104.0989
8.1429 | -104.0726
8.1466 | -104.0491
8.1503 | -104.0252
8.154 | -104.0088
8.1577 | -103.9965
8.1614 | -103.9852
8.1652 | -103.9751
8.1689 | -103.9635
8.1726 | -103.9557
8.1763 | -103.9443
8.18 | -103.9353
8.1837 | -103.923
8.1874 | -103.9087
8.1911 | -103.8943
8.1949 | -103.8795
8.1986 | -103.8644
8.2023 | -103.8444
8.206 | -103.8271
8.2097 | -103.8173
8.2134 | -103.8127
8.2171 | -103.8113
8.2208 | -103.8203
8.2245 | -103.8252
8.2283 | -103.8319
8.232 | -103.8414
8.2357 | -103.847
8.2394 | -103.8544
8.2431 | -103.8616
8.2468 | -103.8673
8.2505 | -103.8783
8.2542 | -103.8843
8.2579 | -103.8899
8.2617 | -103.893
8.2654 | -103.8987
8.2691 | -103.9055
8.2728 | -103.9167
8.2765 | -103.9272
8.2802 | -103.9437
8.2839 | -103.9642
8.2876 | -103.9909
8.2913 | -104.0197
8.2951 | -104.0557
8.2988 | -104.0986
8.3025 | -104.1383
8.3062 | -104.1835
8.3099 | -104.2249
8.3136 | -104.2624
8.3173 | -104.3031
8.321 | -104.3334
8.3248 | -104.3548
8.3285 | -104.3649
8.3322 | -104.3624
8.3359 | -104.3549
8.3396 | -104.3435
8.3433 | -104.3245
8.347 | -104.3094
8.3507 | -104.2912
8.3544 | -104.2762
8.3582 | -104.2723
8.3619 | -104.2655
8.3656 | -104.258
8.3693 | -104.2538
8.373 | -104.2503
8.3767 | -104.2527
8.3804 | -104.2587
8.3841 | -104.2675
8.3878 | -104.2717
8.3916 | -104.2685
8.3953 | -104.2633
8.399 | -104.2598
8.4027 | -104.2519
8.4064 | -104.2378
8.4101 | -104.2234
8.4138 | -104.2123
8.4175 | -104.1977
8.4212 | -104.1809
8.425 | -104.1687
8.4287 | -104.1548
8.4324 | -104.1435
8.4361 | -104.1305
8.4398 | -104.1109
8.4435 | -104.0962
8.4472 | -104.0856
8.4509 | -104.0712
8.4547 | -104.0568
8.4584 | -104.0446
8.4621 | -104.0363
8.4658 | -104.0367
8.4695 | -104.0389
8.4732 | -104.0472
8.4769 | -104.0555
8.4806 | -104.0649
8.4843 | -104.079
8.4881 | -104.0953
8.4918 | -104.1147
8.4955 | -104.1312
8.4992 | -104.1455
8.5029 | -104.1569
8.5066 | -104.1758
8.5103 | -104.1931
8.514 | -104.2109
8.5177 | -104.2215
8.5215 | -104.2306
8.5252 | -104.2458
8.5289 | -104.2649
8.5326 | -104.2776
8.5363 | -104.2911
8.54 | -104.305
8.5437 | -104.3224
8.5474 | -104.3388
8.5511 | -104.3524
8.5549 | -104.3671
8.5586 | -104.3818
8.5623 | -104.401
8.566 | -104.4161
8.5697 | -104.4318
8.5734 | -104.448
8.5771 | -104.463
8.5808 | -104.4746
8.5846 | -104.4852
8.5883 | -104.4926
8.592 | -104.4902
8.5957 | -104.488
8.5994 | -104.4838
8.6031 | -104.4768
8.6068 | -104.4684
8.6105 | -104.4541
8.6142 | -104.4357
8.618 | -104.4136
8.6217 | -104.3896
8.6254 | -104.3643
8.6291 | -104.3405
8.6328 | -104.3144
8.6365 | -104.2929
8.6402 | -104.2755
8.6439 | -104.2647
8.6476 | -104.2502
8.6514 | -104.2351
8.6551 | -104.2217
8.6588 | -104.2125
8.6625 | -104.2045
8.6662 | -104.1983
8.6699 | -104.1964
8.6736 | -104.1997
8.6773 | -104.2048
8.681 | -104.2137
8.6848 | -104.2272
8.6885 | -104.2396
8.6922 | -104.2515
8.6959 | -104.2675
8.6996 | -104.2853
8.7033 | -104.3024
8.707 | -104.325
8.7107 | -104.3494
8.7145 | -104.3724
8.7182 | -104.3996
8.7219 | -104.4306
8.7256 | -104.4634
8.7293 | -104.4928
8.733 | -104.5249
8.7367 | -104.5594
8.7404 | -104.6009
8.7441 | -104.645
8.7479 | -104.6858
8.7516 | -104.7233
8.7553 | -104.7548
8.759 | -104.7841
8.7627 | -104.806
8.7664 | -104.8258
8.7701 | -104.8419
8.7738 | -104.8505
8.7775 | -104.8594
8.7813 | -104.8695
8.785 | -104.8794
8.7887 | -104.8875
8.7924 | -104.8958
8.7961 | -104.9063
8.7998 | -104.9127
8.8035 | -104.9162
8.8072 | -104.9201
8.8109 | -104.9221
8.8147 | -104.921
8.8184 | -104.9173
8.8221 | -104.9172
8.8258 | -104.9046
8.8295 | -104.8878
8.8332 | -104.8714
8.8369 | -104.8534
8.8406 | -104.8365
8.8444 | -104.8187
8.8481 | -104.8019
8.8518 | -104.7851
8.8555 | -104.7655
8.8592 | -104.7486
8.8629 | -104.7279
8.8666 | -104.7026
8.8703 | -104.6762
8.874 | -104.6545
8.8778 | -104.6339
8.8815 | -104.6119
8.8852 | -104.5882
8.8889 | -104.5655
8.8926 | -104.547
8.8963 | -104.5289
8.9 | -104.5131
8.9037 | -104.5021
8.9074 | -104.4937
8.9112 | -104.4833
8.9149 | -104.4717
8.9186 | -104.4596
8.9223 | -104.447
8.926 | -104.4326
8.9297 | -104.422
8.9334 | -104.4134
8.9371 | -104.4039
8.9408 | -104.3919
8.9446 | -104.3853
8.9483 | -104.3793
8.952 | -104.3699
8.9557 | -104.3687
8.9594 | -104.3711
8.9631 | -104.3706
8.9668 | -104.3702
8.9705 | -104.3731
8.9743 | -104.3781
8.978 | -104.3822
8.9817 | -104.3868
8.9854 | -104.3938
8.9891 | -104.4021
8.9928 | -104.4143
8.9965 | -104.4247
9.0002 | -104.4329
9.0039 | -104.4351
9.0077 | -104.4435
9.0114 | -104.452
9.0151 | -104.4571
9.0188 | -104.4578
9.0225 | -104.4562
9.0262 | -104.4507
9.0299 | -104.4403
9.0336 | -104.4305
9.0373 | -104.419
9.0411 | -104.4057
9.0448 | -104.3951
9.0485 | -104.387
9.0522 | -104.3759
9.0559 | -104.3669
9.0596 | -104.3562
9.0633 | -104.3455
9.067 | -104.3398
9.0707 | -104.341
9.0745 | -104.3472
9.0782 | -104.3583
9.0819 | -104.3748
9.0856 | -104.3936
9.0893 | -104.4138
9.093 | -104.4389
9.0967 | -104.4668
9.1004 | -104.4908
9.1042 | -104.5104
9.1079 | -104.5279
9.1116 | -104.5434
9.1153 | -104.5583
9.119 | -104.5715
9.1227 | -104.5853
9.1264 | -104.5956
9.1301 | -104.6069
9.1338 | -104.6177
9.1376 | -104.6292
9.1413 | -104.6383
9.145 | -104.6505
9.1487 | -104.6667
9.1524 | -104.6798
9.1561 | -104.6879
9.1598 | -104.6969
9.1635 | -104.7028
9.1672 | -104.7052
9.171 | -104.7052
9.1747 | -104.7037
9.1784 | -104.7006
9.1821 | -104.6919
9.1858 | -104.6844
9.1895 | -104.6742
9.1932 | -104.6646
9.1969 | -104.6483
9.2006 | -104.6249
9.2044 | -104.5966
9.2081 | -104.5674
9.2118 | -104.5404
9.2155 | -104.5124
9.2192 | -104.4859
9.2229 | -104.4613
9.2266 | -104.4392
9.2303 | -104.4221
9.2341 | -104.4124
9.2378 | -104.4102
9.2415 | -104.4138
9.2452 | -104.4212
9.2489 | -104.4387
9.2526 | -104.4655
9.2563 | -104.4948
9.26 | -104.5251
9.2637 | -104.5587
9.2675 | -104.594
9.2712 | -104.629
9.2749 | -104.6586
9.2786 | -104.682
9.2823 | -104.7031
9.286 | -104.7187
9.2897 | -104.7297
9.2934 | -104.7345
9.2971 | -104.7309
9.3009 | -104.7226
9.3046 | -104.7062
9.3083 | -104.6845
9.312 | -104.6626
9.3157 | -104.636
9.3194 | -104.611
9.3231 | -104.5844
9.3268 | -104.5641
9.3306 | -104.5469
9.3343 | -104.5345
9.338 | -104.5289
9.3417 | -104.5284
9.3454 | -104.5293
9.3491 | -104.5369
9.3528 | -104.5506
9.3565 | -104.5663
9.3602 | -104.5849
9.364 | -104.6054
9.3677 | -104.6212
9.3714 | -104.6324
9.3751 | -104.6466
9.3788 | -104.6591
9.3825 | -104.6671
9.3862 | -104.664
9.3899 | -104.6531
9.3936 | -104.6363
9.3974 | -104.6152
9.4011 | -104.5936
9.4048 | -104.5702
9.4085 | -104.5419
9.4122 | -104.5173
9.4159 | -104.4921
9.4196 | -104.4642
9.4233 | -104.4371
9.427 | -104.419
9.4308 | -104.4094
9.4345 | -104.4038
9.4382 | -104.3986
9.4419 | -104.3968
9.4456 | -104.4027
9.4493 | -104.4129
9.453 | -104.4288
9.4567 | -104.4466
9.4604 | -104.4692
9.4642 | -104.4913
9.4679 | -104.513
9.4716 | -104.5376
9.4753 | -104.5612
9.479 | -104.5829
9.4827 | -104.6068
9.4864 | -104.6293
9.4901 | -104.6502
9.4939 | -104.6718
9.4976 | -104.6922
9.5013 | -104.7131
9.505 | -104.7312
9.5087 | -104.7432
9.5124 | -104.748
9.5161 | -104.7489
9.5198 | -104.74
9.5235 | -104.7305
9.5273 | -104.7235
9.531 | -104.7169
9.5347 | -104.7063
9.5384 | -104.6919
9.5421 | -104.6745
9.5458 | -104.6599
9.5495 | -104.6445
9.5532 | -104.6287
9.5569 | -104.6157
9.5607 | -104.6038
9.5644 | -104.59
9.5681 | -104.5786
9.5718 | -104.57
9.5755 | -104.5636
9.5792 | -104.5596
9.5829 | -104.5527
9.5866 | -104.5482
9.5904 | -104.544
9.5941 | -104.5423
9.5978 | -104.541
9.6015 | -104.5421
9.6052 | -104.5479
9.6089 | -104.5587
9.6126 | -104.5696
9.6163 | -104.5789
9.62 | -104.5895
9.6238 | -104.605
9.6275 | -104.6254
9.6312 | -104.6473
9.6349 | -104.6798
9.6386 | -104.7046
9.6423 | -104.7303
9.646 | -104.7613
9.6497 | -104.7951
9.6534 | -104.8291
9.6572 | -104.8618
9.6609 | -104.9002
9.6646 | -104.9366
9.6683 | -104.9661
9.672 | -104.9928
9.6757 | -105.0255
9.6794 | -105.0592
9.6831 | -105.0941
9.6868 | -105.1263
9.6906 | -105.1532
9.6943 | -105.1743
9.698 | -105.1912
9.7017 | -105.206
9.7054 | -105.2146
9.7091 | -105.2189
9.7128 | -105.2212
9.7165 | -105.2198
9.7203 | -105.2041
9.724 | -105.1848
9.7277 | -105.1585
9.7314 | -105.1307
9.7351 | -105.0975
9.7388 | -105.0593
9.7425 | -105.0203
9.7462 | -104.9811
9.7499 | -104.9366
9.7537 | -104.8898
9.7574 | -104.8535
9.7611 | -104.8198
9.7648 | -104.7866
9.7685 | -104.7549
9.7722 | -104.7315
9.7759 | -104.7097
9.7796 | -104.694
9.7833 | -104.6868
9.7871 | -104.6832
9.7908 | -104.683
9.7945 | -104.6877
9.7982 | -104.6898
9.8019 | -104.6952
9.8056 | -104.704
9.8093 | -104.7103
9.813 | -104.719
9.8167 | -104.7309
9.8205 | -104.7331
9.8242 | -104.7363
9.8279 | -104.7428
9.8316 | -104.7482
9.8353 | -104.7571
9.839 | -104.7679
9.8427 | -104.7747
9.8464 | -104.7787
9.8502 | -104.7853
9.8539 | -104.7943
9.8576 | -104.804
9.8613 | -104.8162
9.865 | -104.833
9.8687 | -104.8478
9.8724 | -104.8602
9.8761 | -104.8723
9.8798 | -104.8814
9.8836 | -104.8915
9.8873 | -104.9006
9.891 | -104.9087
9.8947 | -104.918
9.8984 | -104.9264
9.9021 | -104.9258
9.9058 | -104.9213
9.9095 | -104.9136
9.9132 | -104.9107
9.917 | -104.9057
9.9207 | -104.9005
9.9244 | -104.8933
9.9281 | -104.8817
9.9318 | -104.8685
9.9355 | -104.8561
9.9392 | -104.8423
9.9429 | -104.8325
9.9466 | -104.8268
9.9504 | -104.8165
9.9541 | -104.8096
9.9578 | -104.8017
9.9615 | -104.7962
9.9652 | -104.7967
9.9689 | -104.801
9.9726 | -104.8044
9.9763 | -104.8108
9.9801 | -104.8195
9.9838 | -104.8313
9.9875 | -104.8465
9.9912 | -104.8618
9.9949 | -104.8795
9.9986 | -104.8953
10.0023 | -104.9069
10.006 | -104.9218
10.0097 | -104.9345
10.0135 | -104.9438
10.0172 | -104.9565
10.0209 | -104.9652
10.0246 | -104.9733
10.0283 | -104.9793
10.032 | -104.9835
10.0357 | -104.9857
10.0394 | -104.9873
10.0431 | -104.9875
10.0469 | -104.9875
10.0506 | -104.9883
10.0543 | -104.9876
10.058 | -104.9845
10.0617 | -104.9808
10.0654 | -104.9813
10.0691 | -104.9805
10.0728 | -104.9811
10.0765 | -104.9879
10.0803 | -104.9966
10.084 | -105.0097
10.0877 | -105.0241
10.0914 | -105.0327
10.0951 | -105.0446
10.0988 | -105.0588
10.1025 | -105.0794
10.1062 | -105.1003
10.11 | -105.1205
10.1137 | -105.1338
10.1174 | -105.1451
10.1211 | -105.157
10.1248 | -105.1625
10.1285 | -105.1659
10.1322 | -105.1654
10.1359 | -105.1574
10.1396 | -105.1379
10.1434 | -105.1181
10.1471 | -105.1001
10.1508 | -105.0719
10.1545 | -105.045
10.1582 | -105.0182
10.1619 | -104.9841
10.1656 | -104.952
10.1693 | -104.9233
10.173 | -104.8962
10.1768 | -104.8682
10.1805 | -104.8409
10.1842 | -104.8189
10.1879 | -104.8
10.1916 | -104.7822
10.1953 | -104.7641
10.199 | -104.7494
10.2027 | -104.7377
10.2064 | -104.7263
10.2102 | -104.7216
10.2139 | -104.7203
10.2176 | -104.7218
10.2213 | -104.7261
10.225 | -104.733
10.2287 | -104.74
10.2324 | -104.7467
10.2361 | -104.7569
10.2399 | -104.7735
10.2436 | -104.793
10.2473 | -104.8085
10.251 | -104.822
10.2547 | -104.8416
10.2584 | -104.862
10.2621 | -104.8841
10.2658 | -104.9049
10.2695 | -104.9222
10.2733 | -104.9438
10.277 | -104.9661
10.2807 | -104.9878
10.2844 | -105.0128
10.2881 | -105.0355
10.2918 | -105.0518
10.2955 | -105.0664
10.2992 | -105.0779
10.3029 | -105.0858
10.3067 | -105.0898
10.3104 | -105.093
10.3141 | -105.0935
10.3178 | -105.0868
10.3215 | -105.0791
10.3252 | -105.0667
10.3289 | -105.0523
10.3326 | -105.0351
10.3363 | -105.0164
10.3401 | -104.9993
10.3438 | -104.9865
10.3475 | -104.9766
10.3512 | -104.9639
10.3549 | -104.9539
10.3586 | -104.9445
10.3623 | -104.9291
10.366 | -104.9178
10.3698 | -104.9134
10.3735 | -104.9136
10.3772 | -104.9152
10.3809 | -104.9234
10.3846 | -104.935
10.3883 | -104.9493
10.392 | -104.9638
10.3957 | -104.9784
10.3994 | -104.9967
10.4032 | -105.0158
10.4069 | -105.0287
10.4106 | -105.0383
10.4143 | -105.0461
10.418 | -105.0565
10.4217 | -105.0644
10.4254 | -105.0655
10.4291 | -105.0646
10.4328 | -105.0627
10.4366 | -105.0593
10.4403 | -105.0559
10.444 | -105.0525
10.4477 | -105.0493
10.4514 | -105.0512
10.4551 | -105.0523
10.4588 | -105.0547
10.4625 | -105.0587
10.4662 | -105.0617
10.47 | -105.0674
10.4737 | -105.0751
10.4774 | -105.0863
10.4811 | -105.0995
10.4848 | -105.1144
10.4885 | -105.1271
10.4922 | -105.1399
10.4959 | -105.154
10.4997 | -105.17
10.5034 | -105.1866
10.5071 | -105.2049
10.5108 | -105.2224
10.5145 | -105.2423
10.5182 | -105.2619
10.5219 | -105.2749
10.5256 | -105.2815
10.5293 | -105.2865
10.5331 | -105.2856
10.5368 | -105.2851
10.5405 | -105.2793
10.5442 | -105.2723
10.5479 | -105.2642
10.5516 | -105.2531
10.5553 | -105.2377
10.559 | -105.2204
10.5627 | -105.2
10.5665 | -105.179
10.5702 | -105.1565
10.5739 | -105.1306
10.5776 | -105.1027
10.5813 | -105.0754
10.585 | -105.0537
10.5887 | -105.0322
10.5924 | -105.0122
10.5961 | -104.9963
10.5999 | -104.9792
10.6036 | -104.9587
10.6073 | -104.9399
10.611 | -104.9259
10.6147 | -104.917
10.6184 | -104.9108
10.6221 | -104.9058
10.6258 | -104.9026
10.6296 | -104.8979
10.6333 | -104.8939
10.637 | -104.8953
10.6407 | -104.8979
10.6444 | -104.9059
10.6481 | -104.9164
10.6518 | -104.9271
10.6555 | -104.9426
10.6592 | -104.9625
10.663 | -104.9864
10.6667 | -105.012
10.6704 | -105.0365
10.6741 | -105.0669
10.6778 | -105.1026
10.6815 | -105.136
10.6852 | -105.1691
10.6889 | -105.1984
10.6926 | -105.2283
10.6964 | -105.2549
10.7001 | -105.2802
10.7038 | -105.307
10.7075 | -105.3263
10.7112 | -105.3495
10.7149 | -105.369
10.7186 | -105.3858
10.7223 | -105.3984
10.7261 | -105.4033
10.7298 | -105.408
10.7335 | -105.4116
10.7372 | -105.4093
10.7409 | -105.3995
10.7446 | -105.3898
10.7483 | -105.379
10.752 | -105.3625
10.7557 | -105.3448
10.7595 | -105.3269
10.7632 | -105.3051
10.7669 | -105.2822
10.7706 | -105.2574
10.7743 | -105.2341
10.778 | -105.215
10.7817 | -105.1933
10.7854 | -105.1743
10.7891 | -105.1604
10.7929 | -105.1482
10.7966 | -105.1398
10.8003 | -105.1287
10.804 | -105.1214
10.8077 | -105.1215
10.8114 | -105.1234
10.8151 | -105.1225
10.8188 | -105.1251
10.8225 | -105.133
10.8263 | -105.1463
10.83 | -105.1596
10.8337 | -105.1742
10.8374 | -105.1925
10.8411 | -105.2131
10.8448 | -105.2371
10.8485 | -105.262
10.8522 | -105.2893
10.8559 | -105.3162
10.8597 | -105.3417
10.8634 | -105.3686
10.8671 | -105.399
10.8708 | -105.4286
10.8745 | -105.4558
10.8782 | -105.4797
10.8819 | -105.4967
10.8856 | -105.5164
10.8894 | -105.532
10.8931 | -105.5455
10.8968 | -105.554
10.9005 | -105.5539
10.9042 | -105.547
10.9079 | -105.5393
10.9116 | -105.529
10.9153 | -105.5168
10.919 | -105.5048
10.9228 | -105.4916
10.9265 | -105.4695
10.9302 | -105.4466
10.9339 | -105.4276
10.9376 | -105.4087
10.9413 | -105.3924
10.945 | -105.378
10.9487 | -105.3594
10.9524 | -105.3408
10.9562 | -105.327
10.9599 | -105.3185
10.9636 | -105.3119
10.9673 | -105.3063
10.971 | -105.3018
10.9747 | -105.3043
10.9784 | -105.3077
10.9821 | -105.3105
10.9858 | -105.3125
10.9896 | -105.3158
10.9933 | -105.3217
10.997 | -105.3257
11.0007 | -105.3268
11.0044 | -105.3245
11.0081 | -105.3202
11.0118 | -105.3121
11.0155 | -105.3007
11.0193 | -105.2866
11.023 | -105.2723
11.0267 | -105.252
11.0304 | -105.2335
11.0341 | -105.2174
11.0378 | -105.2024
11.0415 | -105.1867
11.0452 | -105.1675
11.0489 | -105.1507
11.0527 | -105.1354
11.0564 | -105.1259
11.0601 | -105.1227
11.0638 | -105.1204
11.0675 | -105.1191
11.0712 | -105.1208
11.0749 | -105.1271
11.0786 | -105.1346
11.0823 | -105.1464
11.0861 | -105.1659
11.0898 | -105.188
11.0935 | -105.2104
11.0972 | -105.2356
11.1009 | -105.2624
11.1046 | -105.2922
11.1083 | -105.3282
11.112 | -105.3604
11.1158 | -105.3893
11.1195 | -105.4177
11.1232 | -105.4443
11.1269 | -105.4665
11.1306 | -105.4873
11.1343 | -105.4997
11.138 | -105.5068
11.1417 | -105.5053
11.1454 | -105.4988
11.1492 | -105.4888
11.1529 | -105.4739
11.1566 | -105.4553
11.1603 | -105.4356
11.164 | -105.4114
11.1677 | -105.3849
11.1714 | -105.3608
11.1751 | -105.3366
11.1788 | -105.3156
11.1826 | -105.298
11.1863 | -105.2865
11.19 | -105.2803
11.1937 | -105.2792
11.1974 | -105.2835
11.2011 | -105.2939
11.2048 | -105.3036
11.2085 | -105.3187
11.2122 | -105.3398
11.216 | -105.3647
11.2197 | -105.3912
11.2234 | -105.4191
11.2271 | -105.4459
11.2308 | -105.4705
11.2345 | -105.4927
11.2382 | -105.5136
11.2419 | -105.5118
11.2457 | -105.5191
11.2494 | -105.5228
11.2531 | -105.5261
11.2568 | -105.5243
11.2605 | -105.5194
11.2642 | -105.513
11.2679 | -105.5029
11.2716 | -105.4851
11.2753 | -105.4722
11.2791 | -105.4577
11.2828 | -105.4375
11.2865 | -105.4155
11.2902 | -105.3928
11.2939 | -105.372
11.2976 | -105.3547
11.3013 | -105.3389
11.305 | -105.325
11.3087 | -105.3129
11.3125 | -105.2958
11.3162 | -105.2765
11.3199 | -105.262
11.3236 | -105.2507
11.3273 | -105.2414
11.331 | -105.2336
11.3347 | -105.2293
11.3384 | -105.2282
11.3421 | -105.2291
11.3459 | -105.2317
11.3496 | -105.2362
11.3533 | -105.2418
11.357 | -105.2499
11.3607 | -105.259
11.3644 | -105.2709
11.3681 | -105.2893
11.3718 | -105.3033
11.3756 | -105.3148
11.3793 | -105.3278
11.383 | -105.3407
11.3867 | -105.3552
11.3904 | -105.3667
11.3941 | -105.3798
11.3978 | -105.3965
11.4015 | -105.4159
11.4052 | -105.4363
11.409 | -105.4594
11.4127 | -105.4849
11.4164 | -105.5075
11.4201 | -105.525
11.4238 | -105.5477
11.4275 | -105.5703
11.4312 | -105.5944
11.4349 | -105.6143
11.4386 | -105.634
11.4424 | -105.644
11.4461 | -105.6486
11.4498 | -105.6541
11.4535 | -105.6532
11.4572 | -105.6465
11.4609 | -105.6339
11.4646 | -105.6142
11.4683 | -105.5898
11.472 | -105.5622
11.4758 | -105.5285
11.4795 | -105.4953
11.4832 | -105.4593
11.4869 | -105.422
11.4906 | -105.3821
11.4943 | -105.3434
11.498 | -105.3045
11.5017 | -105.2631
11.5055 | -105.2256
11.5092 | -105.1933
11.5129 | -105.166
11.5166 | -105.142
11.5203 | -105.1205
11.524 | -105.1054
11.5277 | -105.0997
11.5314 | -105.098
11.5351 | -105.0977
11.5389 | -105.105
11.5426 | -105.1134
11.5463 | -105.1276
11.55 | -105.1476
11.5537 | -105.1701
11.5574 | -105.1935
11.5611 | -105.2175
11.5648 | -105.2397
11.5685 | -105.2591
11.5723 | -105.2767
11.576 | -105.2951
11.5797 | -105.3104
11.5834 | -105.3265
11.5871 | -105.3449
11.5908 | -105.3575
11.5945 | -105.3685
11.5982 | -105.3755
11.6019 | -105.3862
11.6057 | -105.3953
11.6094 | -105.4038
11.6131 | -105.4139
11.6168 | -105.4225
11.6205 | -105.4333
11.6242 | -105.4436
11.6279 | -105.4531
11.6316 | -105.4627
11.6354 | -105.4746
11.6391 | -105.4884
11.6428 | -105.5036
11.6465 | -105.52
11.6502 | -105.5395
11.6539 | -105.5561
11.6576 | -105.5729
11.6613 | -105.5914
11.665 | -105.611
11.6688 | -105.6354
11.6725 | -105.662
11.6762 | -105.6872
11.6799 | -105.7108
11.6836 | -105.7381
11.6873 | -105.7671
11.691 | -105.7915
11.6947 | -105.8163
11.6984 | -105.8429
11.7022 | -105.8717
11.7059 | -105.8981
11.7096 | -105.9233
11.7133 | -105.9406
11.717 | -105.9528
11.7207 | -105.9608
11.7244 | -105.9712
11.7281 | -105.9783
11.7318 | -105.9822
11.7356 | -105.9821
11.7393 | -105.9794
11.743 | -105.9734
11.7467 | -105.9667
11.7504 | -105.9558
11.7541 | -105.9387
11.7578 | -105.9208
11.7615 | -105.9051
11.7653 | -105.891
11.769 | -105.8699
11.7727 | -105.8522
11.7764 | -105.8429
11.7801 | -105.8338
11.7838 | -105.8214
11.7875 | -105.8065
11.7912 | -105.7931
11.7949 | -105.7823
11.7987 | -105.7725
11.8024 | -105.7658
11.8061 | -105.7579
11.8098 | -105.7487
11.8135 | -105.7396
11.8172 | -105.7302
11.8209 | -105.7194
11.8246 | -105.7129
11.8283 | -105.7063
11.8321 | -105.6997
11.8358 | -105.6959
11.8395 | -105.6905
11.8432 | -105.6806
11.8469 | -105.6729
11.8506 | -105.6685
11.8543 | -105.6619
11.858 | -105.6563
11.8617 | -105.655
11.8655 | -105.6504
11.8692 | -105.6443
11.8729 | -105.6389
11.8766 | -105.6372
11.8803 | -105.6376
11.884 | -105.6365
11.8877 | -105.6371
11.8914 | -105.6374
11.8952 | -105.6361
11.8989 | -105.6338
11.9026 | -105.6297
11.9063 | -105.6302
11.91 | -105.6329
11.9137 | -105.6325
11.9174 | -105.6298
11.9211 | -105.6306
11.9248 | -105.6354
11.9286 | -105.639
11.9323 | -105.6405
11.936 | -105.6366
11.9397 | -105.6351
11.9434 | -105.6334
11.9471 | -105.6257
11.9508 | -105.6133
11.9545 | -105.5966
11.9582 | -105.5823
11.962 | -105.5694
11.9657 | -105.5558
11.9694 | -105.5391
11.9731 | -105.5225
11.9768 | -105.504
11.9805 | -105.4825
11.9842 | -105.4631
11.9879 | -105.4428
11.9916 | -105.4214
11.9954 | -105.3998
11.9991 | -105.3759
12.0028 | -105.3553
12.0065 | -105.3367
12.0102 | -105.3195
12.0139 | -105.3041
12.0176 | -105.292
12.0213 | -105.2811
12.0251 | -105.271
12.0288 | -105.264
12.0325 | -105.2569
12.0362 | -105.2486
12.0399 | -105.2457
12.0436 | -105.2437
12.0473 | -105.2443
12.051 | -105.2435
12.0547 | -105.2445
12.0585 | -105.2495
12.0622 | -105.2553
12.0659 | -105.2605
12.0696 | -105.2666
12.0733 | -105.275
12.077 | -105.283
12.0807 | -105.2591
12.0844 | -105.214
12.0881 | -105.1955
12.0919 | -105.1786
12.0956 | -105.1559
12.0993 | -105.1269
12.103 | -105.0948
12.1067 | -105.0566
12.1104 | -105.0155
12.1141 | -104.9684
12.1178 | -104.9114
12.1215 | -104.8478
12.1253 | -104.7841
12.129 | -104.72
12.1327 | -104.653
12.1364 | -104.5827
12.1401 | -104.5136
12.1438 | -104.4524
12.1475 | -104.3934
12.1512 | -104.3395
12.155 | -104.2988
12.1587 | -104.2652
12.1624 | -104.2432
12.1661 | -104.2336
12.1698 | -104.2342
12.1735 | -104.2458
12.1772 | -104.2671
12.1809 | -104.3009
12.1846 | -104.3449
12.1884 | -104.3946
12.1921 | -104.4535
12.1958 | -104.5227
12.1995 | -104.5988
12.2032 | -104.6787
12.2069 | -104.7622
12.2106 | -104.8446
12.2143 | -104.9229
12.218 | -105.0027
12.2218 | -105.0796
12.2255 | -105.1532
12.2292 | -105.2222
12.2329 | -105.2848
12.2366 | -105.3428
12.2403 | -105.3962
12.244 | -105.4423
12.2477 | -105.4822
12.2515 | -105.5177
12.2552 | -105.5757
12.2589 | -105.6321
12.2626 | -105.6413
12.2663 | -105.6455
12.27 | -105.6473
12.2737 | -105.6486
12.2774 | -105.6501
12.2811 | -105.6495
12.2849 | -105.649
12.2886 | -105.6493
12.2923 | -105.6541
12.296 | -105.6568
12.2997 | -105.6545
12.3034 | -105.6521
12.3071 | -105.6497
12.3108 | -105.6498
12.3145 | -105.6543
12.3183 | -105.6565
12.322 | -105.6562
12.3257 | -105.6604
12.3294 | -105.6678
12.3331 | -105.6731
12.3368 | -105.6809
12.3405 | -105.6883
12.3442 | -105.6949
12.3479 | -105.706
12.3517 | -105.7196
12.3554 | -105.7334
12.3591 | -105.7456
12.3628 | -105.7625
12.3665 | -105.7821
12.3702 | -105.7994
12.3739 | -105.8129
12.3776 | -105.8256
12.3813 | -105.8397
12.3851 | -105.8558
12.3888 | -105.8728
12.3925 | -105.8878
12.3962 | -105.9011
12.3999 | -105.9134
12.4036 | -105.9241
12.4073 | -105.9319
12.411 | -105.938
12.4148 | -105.9418
12.4185 | -105.944
12.4222 | -105.9491
12.4259 | -105.9531
12.4296 | -105.9564
12.4333 | -105.9572
12.437 | -105.9573
12.4407 | -105.9583
12.4444 | -105.9574
12.4482 | -105.9584
12.4519 | -105.9582
12.4556 | -105.9564
12.4593 | -105.9518
12.463 | -105.9451
12.4667 | -105.9346
12.4704 | -105.9229
12.4741 | -105.9127
12.4778 | -105.9021
12.4816 | -105.8863
12.4853 | -105.8682
12.489 | -105.8478
12.4927 | -105.8249
12.4964 | -105.8012
12.5001 | -105.7793
12.5038 | -105.7566
12.5075 | -105.7318
12.5113 | -105.7037
12.515 | -105.6763
12.5187 | -105.6533
12.5224 | -105.6324
12.5261 | -105.6109
12.5298 | -105.5855
12.5335 | -105.5617
12.5372 | -105.5424
12.5409 | -105.5284
12.5447 | -105.5137
12.5484 | -105.5029
12.5521 | -105.4959
12.5558 | -105.4906
12.5595 | -105.4882
12.5632 | -105.4884
12.5669 | -105.4913
12.5706 | -105.4966
12.5743 | -105.5048
12.5781 | -105.5159
12.5818 | -105.5282
12.5855 | -105.5422
12.5892 | -105.5618
12.5929 | -105.5843
12.5966 | -105.608
12.6003 | -105.634
12.604 | -105.6581
12.6077 | -105.684
12.6115 | -105.7112
12.6152 | -105.7358
12.6189 | -105.7619
12.6226 | -105.7855
12.6263 | -105.807
12.63 | -105.8255
12.6337 | -105.8431
12.6374 | -105.8599
12.6412 | -105.8757
12.6449 | -105.8871
12.6486 | -105.8941
12.6523 | -105.8962
12.656 | -105.8936
12.6597 | -105.8881
12.6634 | -105.8825
12.6671 | -105.8756
12.6708 | -105.8645
12.6746 | -105.8501
12.6783 | -105.834
12.682 | -105.8137
12.6857 | -105.7919
12.6894 | -105.7702
12.6931 | -105.7494
12.6968 | -105.7249
12.7005 | -105.7003
12.7042 | -105.6755
12.708 | -105.6518
12.7117 | -105.6343
12.7154 | -105.616
12.7191 | -105.5971
12.7228 | -105.5797
12.7265 | -105.5645
12.7302 | -105.5501
12.7339 | -105.5382
12.7376 | -105.5285
12.7414 | -105.5186
12.7451 | -105.513
12.7488 | -105.5101
12.7525 | -105.5064
12.7562 | -105.5053
12.7599 | -105.504
12.7636 | -105.5078
12.7673 | -105.5133
12.7711 | -105.517
12.7748 | -105.5236
12.7785 | -105.5335
12.7822 | -105.5448
12.7859 | -105.5553
12.7896 | -105.5664
12.7933 | -105.579
12.797 | -105.5938
12.8007 | -105.6064
12.8045 | -105.6219
12.8082 | -105.637
12.8119 | -105.6503
12.8156 | -105.6664
12.8193 | -105.6802
12.823 | -105.6961
12.8267 | -105.711
12.8304 | -105.731
12.8341 | -105.7487
12.8379 | -105.768
12.8416 | -105.787
12.8453 | -105.8017
12.849 | -105.8162
12.8527 | -105.8295
12.8564 | -105.8457
12.8601 | -105.8622
12.8638 | -105.8782
12.8675 | -105.8942
12.8713 | -105.9125
12.875 | -105.9296
12.8787 | -105.9438
12.8824 | -105.9547
12.8861 | -105.9625
12.8898 | -105.9683
12.8935 | -105.9706
12.8972 | -105.9712
12.901 | -105.9711
12.9047 | -105.9704
12.9084 | -105.9644
12.9121 | -105.9601
12.9158 | -105.9567
12.9195 | -105.951
12.9232 | -105.9428
12.9269 | -105.9357
12.9306 | -105.9258
12.9344 | -105.9124
12.9381 | -105.9007
12.9418 | -105.8891
12.9455 | -105.8768
12.9492 | -105.8628
12.9529 | -105.8476
12.9566 | -105.8315
12.9603 | -105.8134
12.964 | -105.8014
12.9678 | -105.7898
12.9715 | -105.7808
12.9752 | -105.7756
12.9789 | -105.7731
12.9826 | -105.7712
12.9863 | -105.7708
12.99 | -105.7727
12.9937 | -105.7763
12.9974 | -105.7829
13.0012 | -105.787
13.0049 | -105.7958
13.0086 | -105.8089
13.0123 | -105.8217
13.016 | -105.8303
13.0197 | -105.8415
13.0234 | -105.853
13.0271 | -105.8658
13.0309 | -105.8809
13.0346 | -105.8953
13.0383 | -105.9069
13.042 | -105.9159
13.0457 | -105.9254
13.0494 | -105.9356
13.0531 | -105.9422
13.0568 | -105.9473
13.0605 | -105.9486
13.0643 | -105.9503
13.068 | -105.9564
13.0717 | -105.9606
13.0754 | -105.9604
13.0791 | -105.9612
13.0828 | -105.9608
13.0865 | -105.9551
13.0902 | -105.9469
13.0939 | -105.9396
13.0977 | -105.9327
13.1014 | -105.925
13.1051 | -105.9151
13.1088 | -105.9051
13.1125 | -105.8934
13.1162 | -105.8806
13.1199 | -105.8682
13.1236 | -105.8561
13.1273 | -105.8462
13.1311 | -105.8334
13.1348 | -105.8194
13.1385 | -105.8044
13.1422 | -105.79
13.1459 | -105.7726
13.1496 | -105.7515
13.1533 | -105.7274
13.157 | -105.7048
13.1608 | -105.6812
13.1645 | -105.6523
13.1682 | -105.6218
13.1719 | -105.5907
13.1756 | -105.5614
13.1793 | -105.5287
13.183 | -105.494
13.1867 | -105.4588
13.1904 | -105.428
13.1942 | -105.3942
13.1979 | -105.3595
13.2016 | -105.3308
13.2053 | -105.3038
13.209 | -105.2755
13.2127 | -105.2507
13.2164 | -105.2282
13.2201 | -105.2103
13.2238 | -105.1979
13.2276 | -105.1907
13.2313 | -105.1873
13.235 | -105.1851
13.2387 | -105.1832
13.2424 | -105.1874
13.2461 | -105.1971
13.2498 | -105.2117
13.2535 | -105.2296
13.2572 | -105.2488
13.261 | -105.2718
13.2647 | -105.2992
13.2684 | -105.3281
13.2721 | -105.3579
13.2758 | -105.3915
13.2795 | -105.426
13.2832 | -105.4583
13.2869 | -105.4896
13.2907 | -105.5217
13.2944 | -105.5562
13.2981 | -105.5882
13.3018 | -105.619
13.3055 | -105.6461
13.3092 | -105.6688
13.3129 | -105.6897
13.3166 | -105.7058
13.3203 | -105.7215
13.3241 | -105.7348
13.3278 | -105.7432
13.3315 | -105.7488
13.3352 | -105.7557
13.3389 | -105.7628
13.3426 | -105.7641
13.3463 | -105.7601
13.35 | -105.7548
13.3537 | -105.7538
13.3575 | -105.7544
13.3612 | -105.7528
13.3649 | -105.7497
13.3686 | -105.7479
13.3723 | -105.7475
13.376 | -105.7486
13.3797 | -105.7456
13.3834 | -105.7411
13.3871 | -105.7373
13.3909 | -105.7344
13.3946 | -105.7308
13.3983 | -105.7292
13.402 | -105.7279
13.4057 | -105.7261
13.4094 | -105.7251
13.4131 | -105.7251
13.4168 | -105.7249
13.4206 | -105.7238
13.4243 | -105.7242
13.428 | -105.726
13.4317 | -105.7301
13.4354 | -105.7328
13.4391 | -105.734
13.4428 | -105.7361
13.4465 | -105.7383
13.4502 | -105.7396
13.454 | -105.7446
13.4577 | -105.7514
13.4614 | -105.7566
13.4651 | -105.7612
13.4688 | -105.7619
13.4725 | -105.7659
13.4762 | -105.7699
13.4799 | -105.7743
13.4836 | -105.7771
13.4874 | -105.7801
13.4911 | -105.7826
13.4948 | -105.7851
13.4985 | -105.7883
13.5022 | -105.793
13.5059 | -105.7987
13.5096 | -105.8048
13.5133 | -105.8078
13.517 | -105.8085
13.5208 | -105.813
13.5245 | -105.817
13.5282 | -105.8198
13.5319 | -105.8201
13.5356 | -105.8175
13.5393 | -105.8183
13.543 | -105.8193
13.5467 | -105.818
13.5505 | -105.8143
13.5542 | -105.8124
13.5579 | -105.8064
13.5616 | -105.8001
13.5653 | -105.7956
13.569 | -105.7883
13.5727 | -105.7825
13.5764 | -105.7789
13.5801 | -105.7749
13.5839 | -105.7696
13.5876 | -105.7624
13.5913 | -105.7525
13.595 | -105.7431
13.5987 | -105.737
13.6024 | -105.7291
13.6061 | -105.7196
13.6098 | -105.7117
13.6135 | -105.7056
13.6173 | -105.7007
13.621 | -105.6945
13.6247 | -105.6904
13.6284 | -105.6878
13.6321 | -105.6851
13.6358 | -105.6809
13.6395 | -105.6777
13.6432 | -105.6757
13.647 | -105.6746
13.6507 | -105.6742
13.6544 | -105.6696
13.6581 | -105.6682
13.6618 | -105.6714
13.6655 | -105.674
13.6692 | -105.6746
13.6729 | -105.6756
13.6766 | -105.6767
13.6804 | -105.6766
13.6841 | -105.6756
13.6878 | -105.6774
13.6915 | -105.6783
13.6952 | -105.6806
13.6989 | -105.6802
13.7026 | -105.6789
13.7063 | -105.6782
13.71 | -105.6758
13.7138 | -105.6759
13.7175 | -105.6759
13.7212 | -105.6773
13.7249 | -105.6799
13.7286 | -105.6854
13.7323 | -105.6928
13.736 | -105.6983
13.7397 | -105.7029
13.7434 | -105.7063
13.7472 | -105.7111
13.7509 | -105.7177
13.7546 | -105.729
13.7583 | -105.7368
13.762 | -105.7411
13.7657 | -105.7475
13.7694 | -105.7538
13.7731 | -105.7584
13.7768 | -105.7625
13.7806 | -105.7647
13.7843 | -105.7675
13.788 | -105.7744
13.7917 | -105.7778
13.7954 | -105.7812
13.7991 | -105.7865
13.8028 | -105.7906
13.8065 | -105.7926
13.8103 | -105.7974
13.814 | -105.8039
13.8177 | -105.8133
13.8214 | -105.8201
13.8251 | -105.826
13.8288 | -105.8344
13.8325 | -105.8433
13.8362 | -105.8515
13.8399 | -105.8598
13.8437 | -105.8661
13.8474 | -105.8698
13.8511 | -105.8774
13.8548 | -105.8844
13.8585 | -105.89
13.8622 | -105.8961
13.8659 | -105.9016
13.8696 | -105.9098
13.8733 | -105.9163
13.8771 | -105.9238
13.8808 | -105.9322
13.8845 | -105.9397
13.8882 | -105.9498
13.8919 | -105.9556
13.8956 | -105.9585
13.8993 | -105.9625
13.903 | -105.9672
13.9068 | -105.9773
13.9105 | -105.9862
13.9142 | -105.9916
13.9179 | -105.9871
13.9216 | -105.9858
13.9253 | -105.9834
13.929 | -105.9823
13.9327 | -105.9782
13.9364 | -105.9742
13.9402 | -105.9689
13.9439 | -105.9588
13.9476 | -105.9525
13.9513 | -105.9463
13.955 | -105.9368
13.9587 | -105.9185
13.9624 | -105.8847
13.9661 | -105.8573
13.9698 | -105.8342
13.9736 | -105.8134
13.9773 | -105.7922
13.981 | -105.7684
13.9847 | -105.7419
13.9884 | -105.7077
13.9921 | -105.6741
13.9958 | -105.6386
13.9995 | -105.5997
14.0032 | -105.5592
14.007 | -105.5184
14.0107 | -105.4783
14.0144 | -105.4385
14.0181 | -105.3968
14.0218 | -105.3552
14.0255 | -105.3143
14.0292 | -105.2718
14.0329 | -105.232
14.0366 | -105.1968
14.0404 | -105.1651
14.0441 | -105.14
14.0478 | -105.1181
14.0515 | -105.0983
14.0552 | -105.0813
14.0589 | -105.0703
14.0626 | -105.0657
14.0663 | -105.0665
14.0701 | -105.0695
14.0738 | -105.0792
14.0775 | -105.0942
14.0812 | -105.1135
14.0849 | -105.1369
14.0886 | -105.1625
14.0923 | -105.1922
14.096 | -105.2282
14.0997 | -105.2626
14.1035 | -105.2993
14.1072 | -105.3352
14.1109 | -105.3695
14.1146 | -105.4059
14.1183 | -105.4442
14.122 | -105.4834
14.1257 | -105.5187
14.1294 | -105.5521
14.1331 | -105.5821
14.1369 | -105.6128
14.1406 | -105.64
14.1443 | -105.6657
14.148 | -105.6899
14.1517 | -105.7066
14.1554 | -105.7209
14.1591 | -105.7313
14.1628 | -105.7489
14.1666 | -105.774
14.1703 | -105.7813
14.174 | -105.779
14.1777 | -105.7718
14.1814 | -105.7652
14.1851 | -105.7588
14.1888 | -105.7513
14.1925 | -105.7455
14.1962 | -105.742
14.2 | -105.738
14.2037 | -105.7344
14.2074 | -105.7303
14.2111 | -105.7264
14.2148 | -105.7191
14.2185 | -105.7124
14.2222 | -105.7067
14.2259 | -105.7052
14.2296 | -105.7035
14.2334 | -105.7028
14.2371 | -105.7044
14.2408 | -105.706
14.2445 | -105.7056
14.2482 | -105.7058
14.2519 | -105.7084
14.2556 | -105.7119
14.2593 | -105.7153
14.263 | -105.7157
14.2668 | -105.7146
14.2705 | -105.716
14.2742 | -105.722
14.2779 | -105.7273
14.2816 | -105.7311
14.2853 | -105.7343
14.289 | -105.7388
14.2927 | -105.7443
14.2965 | -105.751
14.3002 | -105.7566
14.3039 | -105.7625
14.3076 | -105.7667
14.3113 | -105.7719
14.315 | -105.7777
14.3187 | -105.7818
14.3224 | -105.7877
14.3261 | -105.7906
14.3299 | -105.7928
14.3336 | -105.7943
14.3373 | -105.7985
14.341 | -105.8011
14.3447 | -105.8036
14.3484 | -105.8058
14.3521 | -105.8105
14.3558 | -105.8142
14.3595 | -105.8162
14.3633 | -105.8171
14.367 | -105.8154
14.3707 | -105.8127
14.3744 | -105.8111
14.3781 | -105.8085
14.3818 | -105.807
14.3855 | -105.7993
14.3892 | -105.7927
14.3929 | -105.7855
14.3967 | -105.778
14.4004 | -105.766
14.4041 | -105.7516
14.4078 | -105.7364
14.4115 | -105.7255
14.4152 | -105.7141
14.4189 | -105.702
14.4226 | -105.6896
14.4264 | -105.6765
14.4301 | -105.6614
14.4338 | -105.6477
14.4375 | -105.6342
14.4412 | -105.6182
14.4449 | -105.601
14.4486 | -105.587
14.4523 | -105.573
14.456 | -105.5658
14.4598 | -105.556
14.4635 | -105.5493
14.4672 | -105.5454
14.4709 | -105.5419
14.4746 | -105.5383
14.4783 | -105.5351
14.482 | -105.5322
14.4857 | -105.5303
14.4894 | -105.5289
14.4932 | -105.5304
14.4969 | -105.5364
14.5006 | -105.5448
14.5043 | -105.5563
14.508 | -105.5672
14.5117 | -105.5797
14.5154 | -105.5932
14.5191 | -105.6049
14.5228 | -105.6191
14.5266 | -105.634
14.5303 | -105.6472
14.534 | -105.6641
14.5377 | -105.6806
14.5414 | -105.6928
14.5451 | -105.7036
14.5488 | -105.7144
14.5525 | -105.7249
14.5563 | -105.7322
14.56 | -105.7384
14.5637 | -105.7434
14.5674 | -105.7468
14.5711 | -105.751
14.5748 | -105.756
14.5785 | -105.7588
14.5822 | -105.7601
14.5859 | -105.7603
14.5897 | -105.758
14.5934 | -105.7563
14.5971 | -105.7507
14.6008 | -105.7471
14.6045 | -105.7437
14.6082 | -105.7383
14.6119 | -105.7366
14.6156 | -105.7358
14.6193 | -105.7305
14.6231 | -105.7253
14.6268 | -105.7248
14.6305 | -105.725
14.6342 | -105.7247
14.6379 | -105.7239
14.6416 | -105.7265
14.6453 | -105.7301
14.649 | -105.7336
14.6527 | -105.738
14.6565 | -105.7425
14.6602 | -105.7486
14.6639 | -105.755
14.6676 | -105.7596
14.6713 | -105.7652
14.675 | -105.7698
14.6787 | -105.7748
14.6824 | -105.7816
14.6862 | -105.7899
14.6899 | -105.7961
14.6936 | -105.8006
14.6973 | -105.8058
14.701 | -105.8112
14.7047 | -105.8165
14.7084 | -105.8205
14.7121 | -105.8215
14.7158 | -105.823
14.7196 | -105.8292
14.7233 | -105.8357
14.727 | -105.8428
14.7307 | -105.8518
14.7344 | -105.8581
14.7381 | -105.8658
14.7418 | -105.8741
14.7455 | -105.8795
14.7492 | -105.8868
14.753 | -105.8951
14.7567 | -105.8999
14.7604 | -105.9067
14.7641 | -105.9122
14.7678 | -105.9171
14.7715 | -105.9202
14.7752 | -105.9234
14.7789 | -105.9254
14.7826 | -105.9264
14.7864 | -105.925
14.7901 | -105.9234
14.7938 | -105.9206
14.7975 | -105.9179
14.8012 | -105.9153
14.8049 | -105.9137
14.8086 | -105.9136
14.8123 | -105.9086
14.8161 | -105.903
14.8198 | -105.9015
14.8235 | -105.8977
14.8272 | -105.8937
14.8309 | -105.8919
14.8346 | -105.8912
14.8383 | -105.8896
14.842 | -105.8896
14.8457 | -105.8941
14.8495 | -105.8987
14.8532 | -105.9008
14.8569 | -105.9043
14.8606 | -105.9085
14.8643 | -105.912
14.868 | -105.9168
14.8717 | -105.9216
14.8754 | -105.9284
14.8791 | -105.937
14.8829 | -105.9446
14.8866 | -105.9524
14.8903 | -105.9597
14.894 | -105.9656
14.8977 | -105.9726
14.9014 | -105.9769
14.9051 | -105.9813
14.9088 | -105.9866
14.9125 | -105.9923
14.9163 | -105.9945
14.92 | -105.9915
14.9237 | -105.9907
14.9274 | -105.9865
14.9311 | -105.9787
14.9348 | -105.9711
14.9385 | -105.9644
14.9422 | -105.9555
14.946 | -105.9448
14.9497 | -105.9326
14.9534 | -105.922
14.9571 | -105.9109
14.9608 | -105.8982
14.9645 | -105.8872
14.9682 | -105.8768
14.9719 | -105.8649
14.9756 | -105.8547
14.9794 | -105.8494
14.9831 | -105.8426
14.9868 | -105.8325
14.9905 | -105.7475
14.9942 | -105.7255
14.9979 | -105.7041
15.0016 | -105.6812
15.0053 | -105.6582
15.009 | -105.632
15.0128 | -105.603
15.0165 | -105.5718
15.0202 | -105.5397
15.0239 | -105.5039
15.0276 | -105.4662
15.0313 | -105.4293
15.035 | -105.3899
15.0387 | -105.3492
15.0425 | -105.3067
15.0462 | -105.2642
15.0499 | -105.2204
15.0536 | -105.1761
15.0573 | -105.132
15.061 | -105.0889
15.0647 | -105.0507
15.0684 | -105.0144
15.0721 | -104.9809
15.0759 | -104.9516
15.0796 | -104.9254
15.0833 | -104.9044
15.087 | -104.8895
15.0907 | -104.8775
15.0944 | -104.8712
15.0981 | -104.8716
15.1018 | -104.8758
15.1055 | -104.883
15.1093 | -104.8971
15.113 | -104.9155
15.1167 | -104.9411
15.1204 | -104.9687
15.1241 | -104.9983
15.1278 | -105.0303
15.1315 | -105.0638
15.1352 | -105.0993
15.1389 | -105.1358
15.1427 | -105.1703
15.1464 | -105.2031
15.1501 | -105.2351
15.1538 | -105.263
15.1575 | -105.2896
15.1612 | -105.3148
15.1649 | -105.3343
15.1686 | -105.3513
15.1723 | -105.3657
15.1761 | -105.3775
15.1798 | -105.3864
15.1835 | -105.3909
15.1872 | -105.3954
15.1909 | -105.3971
15.1946 | -105.3965
15.1983 | -105.3954
15.202 | -105.3968
15.2058 | -105.431
15.2095 | -105.4488
15.2132 | -105.4407
15.2169 | -105.4341
15.2206 | -105.4291
15.2243 | -105.4279
15.228 | -105.4314
15.2317 | -105.4386
15.2354 | -105.4482
15.2392 | -105.4599
15.2429 | -105.4752
15.2466 | -105.4913
15.2503 | -105.509
15.254 | -105.5272
15.2577 | -105.5495
15.2614 | -105.5746
15.2651 | -105.5992
15.2688 | -105.6261
15.2726 | -105.6557
15.2763 | -105.6869
15.28 | -105.7174
15.2837 | -105.7464
15.2874 | -105.773
15.2911 | -105.8024
15.2948 | -105.8341
15.2985 | -105.8644
15.3023 | -105.8904
15.306 | -105.912
15.3097 | -105.9329
15.3134 | -105.9501
15.3171 | -105.9674
15.3208 | -105.9842
15.3245 | -105.9998
15.3282 | -106.0113
15.3319 | -106.0187
15.3357 | -106.0232
15.3394 | -106.0263
15.3431 | -106.026
15.3468 | -106.0229
15.3505 | -106.0207
15.3542 | -106.0179
15.3579 | -106.0113
15.3616 | -106.0038
15.3653 | -105.9927
15.3691 | -105.9763
15.3728 | -105.9582
15.3765 | -105.9395
15.3802 | -105.9205
15.3839 | -105.9016
15.3876 | -105.8804
15.3913 | -105.8576
15.395 | -105.8325
15.3987 | -105.8086
15.4025 | -105.7843
15.4062 | -105.759
15.4099 | -105.7344
15.4136 | -105.7126
15.4173 | -105.6925
15.421 | -105.6722
15.4247 | -105.6511
15.4284 | -105.6317
15.4322 | -105.6115
15.4359 | -105.5913
15.4396 | -105.5737
15.4433 | -105.5536
15.447 | -105.5361
15.4507 | -105.5199
15.4544 | -105.5072
15.4581 | -105.4961
15.4618 | -105.4863
15.4656 | -105.4806
15.4693 | -105.4742
15.473 | -105.4726
15.4767 | -105.4707
15.4804 | -105.4682
15.4841 | -105.4689
15.4878 | -105.4713
15.4915 | -105.4728
15.4952 | -105.4776
15.499 | -105.4821
15.5027 | -105.4876
15.5064 | -105.4966
15.5101 | -105.4996
15.5138 | -105.5028
15.5175 | -105.5109
15.5212 | -105.5205
15.5249 | -105.5258
15.5286 | -105.5349
15.5324 | -105.5423
15.5361 | -105.5506
15.5398 | -105.5572
15.5435 | -105.5648
15.5472 | -105.5706
15.5509 | -105.5718
15.5546 | -105.5722
15.5583 | -105.574
15.5621 | -105.5781
15.5658 | -105.582
15.5695 | -105.5842
15.5732 | -105.5849
15.5769 | -105.584
15.5806 | -105.5843
15.5843 | -105.5862
15.588 | -105.5901
15.5917 | -105.5965
15.5955 | -105.6039
15.5992 | -105.6108
15.6029 | -105.6126
15.6066 | -105.6145
15.6103 | -105.6199
15.614 | -105.6265
15.6177 | -105.6339
15.6214 | -105.6424
15.6251 | -105.6518
15.6289 | -105.6612
15.6326 | -105.6696
15.6363 | -105.6778
15.64 | -105.6876
15.6437 | -105.6984
15.6474 | -105.7128
15.6511 | -105.7261
15.6548 | -105.7379
15.6585 | -105.75
15.6623 | -105.7637
15.666 | -105.78
15.6697 | -105.7944
15.6734 | -105.8091
15.6771 | -105.8214
15.6808 | -105.8343
15.6845 | -105.8455
15.6882 | -105.8553
15.692 | -105.8651
15.6957 | -105.8713
15.6994 | -105.8779
15.7031 | -105.8849
15.7068 | -105.89
15.7105 | -105.894
15.7142 | -105.8991
15.7179 | -105.9035
15.7216 | -105.91
15.7254 | -105.9154
15.7291 | -105.9193
15.7328 | -105.9242
15.7365 | -105.9286
15.7402 | -105.9321
15.7439 | -105.936
15.7476 | -105.9392
15.7513 | -105.9403
15.755 | -105.9369
15.7588 | -105.9355
15.7625 | -105.9345
15.7662 | -105.9286
15.7699 | -105.9237
15.7736 | -105.918
15.7773 | -105.9132
15.781 | -105.9079
15.7847 | -105.8984
15.7884 | -105.886
15.7922 | -105.8723
15.7959 | -105.8603
15.7996 | -105.8463
15.8033 | -105.8301
15.807 | -105.8121
15.8107 | -105.7932
15.8144 | -105.7738
15.8181 | -105.7543
15.8219 | -105.7357
15.8256 | -105.718
15.8293 | -105.7048
15.833 | -105.692
15.8367 | -105.674
15.8404 | -105.6578
15.8441 | -105.6425
15.8478 | -105.626
15.8515 | -105.6085
15.8553 | -105.5933
15.859 | -105.5819
15.8627 | -105.57
15.8664 | -105.559
15.8701 | -105.549
15.8738 | -105.5382
15.8775 | -105.5307
15.8812 | -105.5256
15.8849 | -105.5198
15.8887 | -105.5145
15.8924 | -105.5099
15.8961 | -105.5067
15.8998 | -105.5059
15.9035 | -105.5084
15.9072 | -105.5111
15.9109 | -105.5147
15.9146 | -105.517
15.9183 | -105.5178
15.9221 | -105.5224
15.9258 | -105.5288
15.9295 | -105.5382
15.9332 | -105.5475
15.9369 | -105.5556
15.9406 | -105.5627
15.9443 | -105.5687
15.948 | -105.5773
15.9518 | -105.5868
15.9555 | -105.597
15.9592 | -105.6061
15.9629 | -105.6136
15.9666 | -105.6214
15.9703 | -105.6291
15.974 | -105.6386
15.9777 | -105.6504
15.9814 | -105.6646
15.9852 | -105.6754
15.9889 | -105.6818
15.9926 | -105.6921
15.9963 | -105.701
16 | -105.7131
16.0037 | -105.7242
16.0074 | -105.7328
16.0111 | -105.7393
16.0148 | -105.7451
16.0186 | -105.7517
16.0223 | -105.7569
16.026 | -105.7617
16.0297 | -105.768
16.0334 | -105.7746
16.0371 | -105.7795
16.0408 | -105.7834
16.0445 | -105.7867
16.0482 | -105.7887
16.052 | -105.7886
16.0557 | -105.7884
16.0594 | -105.7866
16.0631 | -105.7872
16.0668 | -105.7878
16.0705 | -105.7905
16.0742 | -105.7954
16.0779 | -105.7985
16.0817 | -105.8035
16.0854 | -105.8055
16.0891 | -105.8062
16.0928 | -105.8058
16.0965 | -105.8064
16.1002 | -105.8077
16.1039 | -105.807
16.1076 | -105.8078
16.1113 | -105.8091
16.1151 | -105.8128
16.1188 | -105.8142
16.1225 | -105.8135
16.1262 | -105.8138
16.1299 | -105.8163
16.1336 | -105.8169
16.1373 | -105.8196
16.141 | -105.8215
16.1447 | -105.8244
16.1485 | -105.8313
16.1522 | -105.8359
16.1559 | -105.8378
16.1596 | -105.8409
16.1633 | -105.8471
16.167 | -105.8531
16.1707 | -105.8604
16.1744 | -105.8684
16.1781 | -105.8738
16.1819 | -105.8788
16.1856 | -105.8842
16.1893 | -105.8947
16.193 | -105.9037
16.1967 | -105.9104
16.2004 | -105.9171
16.2041 | -105.9257
16.2078 | -105.9329
16.2116 | -105.9399
16.2153 | -105.9483
16.219 | -105.9568
16.2227 | -105.9618
16.2264 | -105.9664
16.2301 | -105.9693
16.2338 | -105.974
16.2375 | -105.9796
16.2412 | -105.9865
16.245 | -105.9933
16.2487 | -105.9968
16.2524 | -105.999
16.2561 | -106.0001
16.2598 | -106.0015
16.2635 | -106.0026
16.2672 | -106.0016
16.2709 | -105.9994
16.2746 | -105.9961
16.2784 | -105.9924
16.2821 | -105.9895
16.2858 | -105.987
16.2895 | -105.9859
16.2932 | -105.9827
16.2969 | -105.9791
16.3006 | -105.9767
16.3043 | -105.9721
16.308 | -105.9685
16.3118 | -105.9651
16.3155 | -105.9614
16.3192 | -105.9576
16.3229 | -105.955
16.3266 | -105.9534
16.3303 | -105.9514
16.334 | -105.9499
16.3377 | -105.9495
16.3415 | -105.9466
16.3452 | -105.9464
16.3489 | -105.9456
16.3526 | -105.9452
16.3563 | -105.9452
16.36 | -105.9436
16.3637 | -105.9394
16.3674 | -105.9364
16.3711 | -105.9346
16.3749 | -105.9361
16.3786 | -105.9361
16.3823 | -105.9317
16.386 | -105.9265
16.3897 | -105.9228
16.3934 | -105.9175
16.3971 | -105.911
16.4008 | -105.9074
16.4045 | -105.9026
16.4083 | -105.8979
16.412 | -105.8936
16.4157 | -105.8898
16.4194 | -105.8888
16.4231 | -105.8865
16.4268 | -105.8834
16.4305 | -105.8799
16.4342 | -105.8799
16.4379 | -105.879
16.4417 | -105.8784
16.4454 | -105.8771
16.4491 | -105.8779
16.4528 | -105.8839
16.4565 | -105.8892
16.4602 | -105.8938
16.4639 | -105.9004
16.4676 | -105.9075
16.4714 | -105.9146
16.4751 | -105.9231
16.4788 | -105.9297
16.4825 | -105.9412
16.4862 | -105.9555
16.4899 | -105.9717
16.4936 | -105.9862
16.4973 | -105.9981
16.501 | -106.0129
16.5048 | -106.0284
16.5085 | -106.0444
16.5122 | -106.0612
16.5159 | -106.0779
16.5196 | -106.0946
16.5233 | -106.1108
16.527 | -106.1236
16.5307 | -106.1377
16.5344 | -106.1504
16.5382 | -106.1627
16.5419 | -106.1746
16.5456 | -106.1889
16.5493 | -106.2059
16.553 | -106.2233
16.5567 | -106.2357
16.5604 | -106.2472
16.5641 | -106.2584
16.5678 | -106.2693
16.5716 | -106.2791
16.5753 | -106.2896
16.579 | -106.2988
16.5827 | -106.3074
16.5864 | -106.315
16.5901 | -106.3233
16.5938 | -106.3307
16.5975 | -106.3336
16.6013 | -106.3407
16.605 | -106.3476
16.6087 | -106.3522
16.6124 | -106.3559
16.6161 | -106.3585
16.6198 | -106.3602
16.6235 | -106.3618
16.6272 | -106.3616
16.6309 | -106.3626
16.6347 | -106.3633
16.6384 | -106.3602
16.6421 | -106.3591
16.6458 | -106.3602
16.6495 | -106.3621
16.6532 | -106.3652
16.6569 | -106.3651
16.6606 | -106.3633
16.6643 | -106.3628
16.6681 | -106.3605
16.6718 | -106.3579
16.6755 | -106.3582
16.6792 | -106.3567
16.6829 | -106.3539
16.6866 | -106.349
16.6903 | -106.3421
16.694 | -106.3385
16.6977 | -106.3363
16.7015 | -106.3305
16.7052 | -106.324
16.7089 | -106.3169
16.7126 | -106.3096
16.7163 | -106.3054
16.72 | -106.2991
16.7237 | -106.2905
16.7274 | -106.2808
16.7312 | -106.2698
16.7349 | -106.2632
16.7386 | -106.2545
16.7423 | -106.2443
16.746 | -106.2308
16.7497 | -106.2174
16.7534 | -106.205
16.7571 | -106.1917
16.7608 | -106.1766
16.7646 | -106.1636
16.7683 | -106.152
16.772 | -106.1409
16.7757 | -106.1297
16.7794 | -106.1159
16.7831 | -106.1013
16.7868 | -106.0851
16.7905 | -106.0709
16.7942 | -106.0603
16.798 | -106.0494
16.8017 | -106.0403
16.8054 | -106.0299
16.8091 | -106.0213
16.8128 | -106.0135
16.8165 | -106.0069
16.8202 | -106.0016
16.8239 | -105.9974
16.8276 | -105.9957
16.8314 | -105.9958
16.8351 | -105.9953
16.8388 | -105.9971
16.8425 | -105.9994
16.8462 | -106.0036
16.8499 | -106.0096
16.8536 | -106.0167
16.8573 | -106.0263
16.8611 | -106.0369
16.8648 | -106.0489
16.8685 | -106.0622
16.8722 | -106.072
16.8759 | -106.0808
16.8796 | -106.0944
16.8833 | -106.1083
16.887 | -106.1183
16.8907 | -106.1286
16.8945 | -106.1385
16.8982 | -106.1488
16.9019 | -106.1603
16.9056 | -106.174
16.9093 | -106.1895
16.913 | -106.2011
16.9167 | -106.2082
16.9204 | -106.2173
16.9241 | -106.2266
16.9279 | -106.234
16.9316 | -106.237
16.9353 | -106.239
16.939 | -106.2394
16.9427 | -106.2386
16.9464 | -106.2357
16.9501 | -106.2329
16.9538 | -106.2295
16.9576 | -106.2223
16.9613 | -106.2126
16.965 | -106.2022
16.9687 | -106.1938
16.9724 | -106.184
16.9761 | -106.1726
16.9798 | -106.1614
16.9835 | -106.1498
16.9872 | -106.1395
16.991 | -106.1271
16.9947 | -106.113
16.9984 | -106.097
17.0021 | -106.084
17.0058 | -106.0694
17.0095 | -106.0497
17.0132 | -106.0316
17.0169 | -106.0137
17.0206 | -105.996
17.0244 | -105.9779
17.0281 | -105.9591
17.0318 | -105.9405
17.0355 | -105.9213
17.0392 | -105.9033
17.0429 | -105.8849
17.0466 | -105.868
17.0503 | -105.8498
17.054 | -105.8302
17.0578 | -105.8111
17.0615 | -105.7925
17.0652 | -105.7744
17.0689 | -105.7579
17.0726 | -105.7421
17.0763 | -105.7259
17.08 | -105.7103
17.0837 | -105.6975
17.0875 | -105.6849
17.0912 | -105.6722
17.0949 | -105.6594
17.0986 | -105.648
17.1023 | -105.6374
17.106 | -105.6281
17.1097 | -105.6201
17.1134 | -105.6129
17.1171 | -105.6072
17.1209 | -105.6018
17.1246 | -105.597
17.1283 | -105.5952
17.132 | -105.5997
17.1357 | -105.5987
17.1394 | -105.597
17.1431 | -105.5994
17.1468 | -105.6024
17.1505 | -105.6061
17.1543 | -105.6099
17.158 | -105.6151
17.1617 | -105.6235
17.1654 | -105.6305
17.1691 | -105.6388
17.1728 | -105.6498
17.1765 | -105.6637
17.1802 | -105.6765
17.1839 | -105.6889
17.1877 | -105.7036
17.1914 | -105.721
17.1951 | -105.7396
17.1988 | -105.7557
17.2025 | -105.7716
17.2062 | -105.7892
17.2099 | -105.8071
17.2136 | -105.8243
17.2174 | -105.842
17.2211 | -105.859
17.2248 | -105.8742
17.2285 | -105.8919
17.2322 | -105.9084
17.2359 | -105.9224
17.2396 | -105.9369
17.2433 | -105.9523
17.247 | -105.9662
17.2508 | -105.977
17.2545 | -105.9909
17.2582 | -106.0037
17.2619 | -106.0138
17.2656 | -106.0247
17.2693 | -106.0349
17.273 | -106.0437
17.2767 | -106.0496
17.2804 | -106.0547
17.2842 | -106.0571
17.2879 | -106.0578
17.2916 | -106.058
17.2953 | -106.0589
17.299 | -106.0588
17.3027 | -106.0553
17.3064 | -106.0496
17.3101 | -106.0439
17.3138 | -106.0357
17.3176 | -106.028
17.3213 | -106.0215
17.325 | -106.0124
17.3287 | -106.0021
17.3324 | -105.9916
17.3361 | -105.9807
17.3398 | -105.97
17.3435 | -105.9619
17.3473 | -105.9523
17.351 | -105.9414
17.3547 | -105.9297
17.3584 | -105.9217
17.3621 | -105.9168
17.3658 | -105.9102
17.3695 | -105.9047
17.3732 | -105.9028
17.3769 | -105.9042
17.3807 | -105.9036
17.3844 | -105.9029
17.3881 | -105.904
17.3918 | -105.9046
17.3955 | -105.9079
17.3992 | -105.9102
17.4029 | -105.9142
17.4066 | -105.9202
17.4103 | -105.9265
17.4141 | -105.9347
17.4178 | -105.9417
17.4215 | -105.9454
17.4252 | -105.9502
17.4289 | -105.9544
17.4326 | -105.9607
17.4363 | -105.9678
17.44 | -105.9723
17.4437 | -105.9768
17.4475 | -105.9843
17.4512 | -105.9938
17.4549 | -106.0021
17.4586 | -106.0097
17.4623 | -106.016
17.466 | -106.0263
17.4697 | -106.0352
17.4734 | -106.0435
17.4772 | -106.0518
17.4809 | -106.0606
17.4846 | -106.0716
17.4883 | -106.0823
17.492 | -106.0927
17.4957 | -106.102
17.4994 | -106.1132
17.5031 | -106.1232
17.5068 | -106.133
17.5106 | -106.1456
17.5143 | -106.1568
17.518 | -106.1658
17.5217 | -106.178
17.5254 | -106.1895
17.5291 | -106.1987
17.5328 | -106.2063
17.5365 | -106.2137
17.5402 | -106.2221
17.544 | -106.228
17.5477 | -106.232
17.5514 | -106.2349
17.5551 | -106.2358
17.5588 | -106.2338
17.5625 | -106.2325
17.5662 | -106.229
17.5699 | -106.2248
17.5736 | -106.2187
17.5774 | -106.2086
17.5811 | -106.1992
17.5848 | -106.1898
17.5885 | -106.1798
17.5922 | -106.1665
17.5959 | -106.1515
17.5996 | -106.1328
17.6033 | -106.1151
17.6071 | -106.0972
17.6108 | -106.0781
17.6145 | -106.0587
17.6182 | -106.0391
17.6219 | -106.0184
17.6256 | -105.9971
17.6293 | -105.9756
17.633 | -105.9547
17.6367 | -105.9369
17.6405 | -105.9173
17.6442 | -105.8958
17.6479 | -105.8762
17.6516 | -105.8582
17.6553 | -105.8392
17.659 | -105.8205
17.6627 | -105.804
17.6664 | -105.7934
17.6701 | -105.7795
17.6739 | -105.7678
17.6776 | -105.7543
17.6813 | -105.7425
17.685 | -105.7294
17.6887 | -105.7165
17.6924 | -105.7056
17.6961 | -105.6934
17.6998 | -105.683
17.7035 | -105.6735
17.7073 | -105.6646
17.711 | -105.6578
17.7147 | -105.6518
17.7184 | -105.6441
17.7221 | -105.6384
17.7258 | -105.635
17.7295 | -105.6315
17.7332 | -105.6288
17.737 | -105.6259
17.7407 | -105.6239
17.7444 | -105.623
17.7481 | -105.6218
17.7518 | -105.6216
17.7555 | -105.6231
17.7592 | -105.6242
17.7629 | -105.6248
17.7666 | -105.6295
17.7704 | -105.6375
17.7741 | -105.644
17.7778 | -105.6507
17.7815 | -105.6591
17.7852 | -105.671
17.7889 | -105.6842
17.7926 | -105.6964
17.7963 | -105.7095
17.8 | -105.7243
17.8038 | -105.7417
17.8075 | -105.7613
17.8112 | -105.7812
17.8149 | -105.7986
17.8186 | -105.8184
17.8223 | -105.84
17.826 | -105.8626
17.8297 | -105.8866
17.8334 | -105.9081
17.8372 | -105.9299
17.8409 | -105.9536
17.8446 | -105.9781
17.8483 | -106.002
17.852 | -106.0263
17.8557 | -106.0504
17.8594 | -106.0727
17.8631 | -106.0917
17.8669 | -106.1112
17.8706 | -106.1297
17.8743 | -106.143
17.878 | -106.1577
17.8817 | -106.1753
17.8854 | -106.1889
17.8891 | -106.1988
17.8928 | -106.2076
17.8965 | -106.2163
17.9003 | -106.2217
17.904 | -106.2226
17.9077 | -106.2218
17.9114 | -106.2223
17.9151 | -106.2195
17.9188 | -106.2157
17.9225 | -106.2114
17.9262 | -106.2041
17.9299 | -106.1948
17.9337 | -106.1862
17.9374 | -106.1748
17.9411 | -106.1625
17.9448 | -106.1509
17.9485 | -106.1384
17.9522 | -106.1257
17.9559 | -106.1106
17.9596 | -106.0948
17.9633 | -106.0796
17.9671 | -106.0634
17.9708 | -106.049
17.9745 | -106.0366
17.9782 | -106.0247
17.9819 | -106.0146
17.9856 | -106.0043
17.9893 | -105.9957
17.993 | -105.9898
17.9968 | -105.9835
18.0005 | -105.9791
18.0042 | -105.9765
18.0079 | -105.9739
18.0116 | -105.972
18.0153 | -105.9721
18.019 | -105.9756
18.0227 | -105.9771
18.0264 | -105.9804
18.0302 | -105.9887
18.0339 | -106.0009
18.0376 | -106.0131
18.0413 | -106.0255
18.045 | -106.0366
18.0487 | -106.0504
18.0524 | -106.0663
18.0561 | -106.0842
18.0598 | -106.1035
18.0636 | -106.1211
18.0673 | -106.1361
18.071 | -106.1508
18.0747 | -106.1639
18.0784 | -106.1789
18.0821 | -106.1947
18.0858 | -106.2081
18.0895 | -106.2227
18.0933 | -106.2344
18.097 | -106.2439
18.1007 | -106.252
18.1044 | -106.2608
18.1081 | -106.2682
18.1118 | -106.2711
18.1155 | -106.2735
18.1192 | -106.2743
18.1229 | -106.2717
18.1267 | -106.2665
18.1304 | -106.2645
18.1341 | -106.2564
18.1378 | -106.2448
18.1415 | -106.2363
18.1452 | -106.2253
18.1489 | -106.2138
18.1526 | -106.2003
18.1563 | -106.1863
18.1601 | -106.1748
18.1638 | -106.1611
18.1675 | -106.1426
18.1712 | -106.1256
18.1749 | -106.1069
18.1786 | -106.0883
18.1823 | -106.0712
18.186 | -106.0561
18.1897 | -106.0382
18.1935 | -106.0211
18.1972 | -106.0046
18.2009 | -105.9878
18.2046 | -105.9778
18.2083 | -105.9626
18.212 | -105.9484
18.2157 | -105.9356
18.2194 | -105.9247
18.2232 | -105.9142
18.2269 | -105.9059
18.2306 | -105.8984
18.2343 | -105.89
18.238 | -105.8831
18.2417 | -105.8774
18.2454 | -105.8727
18.2491 | -105.8678
18.2528 | -105.8649
18.2566 | -105.863
18.2603 | -105.8605
18.264 | -105.8609
18.2677 | -105.8643
18.2714 | -105.8675
18.2751 | -105.8701
18.2788 | -105.8737
18.2825 | -105.8777
18.2862 | -105.8811
18.29 | -105.8832
18.2937 | -105.8849
18.2974 | -105.8878
18.3011 | -105.8941
18.3048 | -105.9007
18.3085 | -105.9044
18.3122 | -105.909
18.3159 | -105.9108
18.3196 | -105.9093
18.3234 | -105.9122
18.3271 | -105.92
18.3308 | -105.9258
18.3345 | -105.9295
18.3382 | -105.9303
18.3419 | -105.9278
18.3456 | -105.9256
18.3493 | -105.9205
18.353 | -105.9167
18.3568 | -105.912
18.3605 | -105.9079
18.3642 | -105.9043
18.3679 | -105.8994
18.3716 | -105.8941
18.3753 | -105.888
18.379 | -105.8835
18.3827 | -105.879
18.3865 | -105.8736
18.3902 | -105.8664
18.3939 | -105.8611
18.3976 | -105.8554
18.4013 | -105.8493
18.405 | -105.8427
18.4087 | -105.8367
18.4124 | -105.8312
18.4161 | -105.8269
18.4199 | -105.8212
18.4236 | -105.8164
18.4273 | -105.8156
18.431 | -105.8158
18.4347 | -105.8171
18.4384 | -105.8175
18.4421 | -105.8206
18.4458 | -105.8244
18.4495 | -105.8303
18.4533 | -105.8378
18.457 | -105.8455
18.4607 | -105.8526
18.4644 | -105.8608
18.4681 | -105.8705
18.4718 | -105.8804
18.4755 | -105.8928
18.4792 | -105.906
18.483 | -105.9179
18.4867 | -105.9317
18.4904 | -105.9452
18.4941 | -105.9579
18.4978 | -105.9718
18.5015 | -105.9853
18.5052 | -105.997
18.5089 | -106.0113
18.5126 | -106.026
18.5164 | -106.0383
18.5201 | -106.0517
18.5238 | -106.0646
18.5275 | -106.0742
18.5312 | -106.0847
18.5349 | -106.0936
18.5386 | -106.1032
18.5423 | -106.1135
18.546 | -106.1221
18.5498 | -106.1291
18.5535 | -106.1354
18.5572 | -106.1419
18.5609 | -106.1478
18.5646 | -106.1482
18.5683 | -106.1499
18.572 | -106.1538
18.5757 | -106.1552
18.5794 | -106.1567
18.5832 | -106.1585
18.5869 | -106.1574
18.5906 | -106.1532
18.5943 | -106.1501
18.598 | -106.1479
18.6017 | -106.1454
18.6054 | -106.1459
18.6091 | -106.1435
18.6129 | -106.1399
18.6166 | -106.1341
18.6203 | -106.126
18.624 | -106.1151
18.6277 | -106.1039
18.6314 | -106.0951
18.6351 | -106.0855
18.6388 | -106.0747
18.6425 | -106.0655
18.6463 | -106.058
18.65 | -106.0509
18.6537 | -106.0448
18.6574 | -106.0371
18.6611 | -106.0292
18.6648 | -106.0198
18.6685 | -106.0118
18.6722 | -106.0048
18.6759 | -105.9989
18.6797 | -105.9918
18.6834 | -105.9859
18.6871 | -105.9824
18.6908 | -105.9772
18.6945 | -105.9717
18.6982 | -105.9673
18.7019 | -105.9662
18.7056 | -105.9646
18.7093 | -105.9623
18.7131 | -105.9631
18.7168 | -105.9648
18.7205 | -105.966
18.7242 | -105.9685
18.7279 | -105.9692
18.7316 | -105.9704
18.7353 | -105.9767
18.739 | -105.9861
18.7428 | -105.9917
18.7465 | -105.995
18.7502 | -105.9983
18.7539 | -106.0014
18.7576 | -106.0055
18.7613 | -106.0107
18.765 | -106.0147
18.7687 | -106.0167
18.7724 | -106.0204
18.7762 | -106.0222
18.7799 | -106.0225
18.7836 | -106.0235
18.7873 | -106.0226
18.791 | -106.0229
18.7947 | -106.0222
18.7984 | -106.0175
18.8021 | -106.0113
18.8058 | -106.0053
18.8096 | -105.9986
18.8133 | -105.9939
18.817 | -105.9903
18.8207 | -105.984
18.8244 | -105.977
18.8281 | -105.9699
18.8318 | -105.9609
18.8355 | -105.9516
18.8392 | -105.9437
18.843 | -105.935
18.8467 | -105.9242
18.8504 | -105.9137
18.8541 | -105.9047
18.8578 | -105.894
18.8615 | -105.883
18.8652 | -105.8702
18.8689 | -105.8568
18.8727 | -105.8453
18.8764 | -105.8367
18.8801 | -105.8286
18.8838 | -105.8197
18.8875 | -105.8104
18.8912 | -105.8005
18.8949 | -105.7907
18.8986 | -105.7806
18.9023 | -105.7718
18.9061 | -105.7634
18.9098 | -105.7532
18.9135 | -105.7438
18.9172 | -105.7344
18.9209 | -105.7256
18.9246 | -105.7191
18.9283 | -105.7139
18.932 | -105.7113
18.9357 | -105.7087
18.9395 | -105.7054
18.9432 | -105.703
18.9469 | -105.7049
18.9506 | -105.7066
18.9543 | -105.7083
18.958 | -105.7113
18.9617 | -105.7156
18.9654 | -105.7199
18.9691 | -105.7246
18.9729 | -105.7292
18.9766 | -105.7354
18.9803 | -105.7427
18.984 | -105.7497
18.9877 | -105.757
18.9914 | -105.7648
18.9951 | -105.7746
18.9988 | -105.784
19.0026 | -105.7882
19.0063 | -105.7929
19.01 | -105.8003
19.0137 | -105.8067
19.0174 | -105.8161
19.0211 | -105.8258
19.0248 | -105.8336
19.0285 | -105.8402
19.0322 | -105.8455
19.036 | -105.8506
19.0397 | -105.8551
19.0434 | -105.8567
19.0471 | -105.857
19.0508 | -105.8549
19.0545 | -105.851
19.0582 | -105.8488
19.0619 | -105.8424
19.0656 | -105.8362
19.0694 | -105.8316
19.0731 | -105.8272
19.0768 | -105.8186
19.0805 | -105.8077
19.0842 | -105.7959
19.0879 | -105.783
19.0916 | -105.7705
19.0953 | -105.7573
19.099 | -105.7425
19.1028 | -105.7304
19.1065 | -105.7173
19.1102 | -105.7013
19.1139 | -105.6842
19.1176 | -105.6685
19.1213 | -105.6535
19.125 | -105.6383
19.1287 | -105.6262
19.1325 | -105.615
19.1362 | -105.6038
19.1399 | -105.5945
19.1436 | -105.5849
19.1473 | -105.5735
19.151 | -105.5626
19.1547 | -105.5549
19.1584 | -105.5466
19.1621 | -105.5408
19.1659 | -105.5358
19.1696 | -105.5322
19.1733 | -105.5283
19.177 | -105.5236
19.1807 | -105.5206
19.1844 | -105.5161
19.1881 | -105.5128
19.1918 | -105.5118
19.1955 | -105.5105
19.1993 | -105.5086
19.203 | -105.5053
19.2067 | -105.5016
19.2104 | -105.4984
19.2141 | -105.4971
19.2178 | -105.4968
19.2215 | -105.4982
19.2252 | -105.5
19.2289 | -105.501
19.2327 | -105.5025
19.2364 | -105.5058
19.2401 | -105.5081
19.2438 | -105.5107
19.2475 | -105.514
19.2512 | -105.5174
19.2549 | -105.5217
19.2586 | -105.5264
19.2624 | -105.5309
19.2661 | -105.5351
19.2698 | -105.5403
19.2735 | -105.5503
19.2772 | -105.5572
19.2809 | -105.5656
19.2846 | -105.5723
19.2883 | -105.5799
19.292 | -105.5873
19.2958 | -105.595
19.2995 | -105.6035
19.3032 | -105.6111
19.3069 | -105.6208
19.3106 | -105.6318
19.3143 | -105.6427
19.318 | -105.6537
19.3217 | -105.6662
19.3254 | -105.6806
19.3292 | -105.6945
19.3329 | -105.708
19.3366 | -105.7228
19.3403 | -105.7352
19.344 | -105.7452
19.3477 | -105.7579
19.3514 | -105.7718
19.3551 | -105.7841
19.3588 | -105.795
19.3626 | -105.806
19.3663 | -105.8174
19.37 | -105.8266
19.3737 | -105.8356
19.3774 | -105.8431
19.3811 | -105.8489
19.3848 | -105.8552
19.3885 | -105.8651
19.3923 | -105.8749
19.396 | -105.8807
19.3997 | -105.8827
19.4034 | -105.8832
19.4071 | -105.8804
19.4108 | -105.8756
19.4145 | -105.8684
19.4182 | -105.8599
19.4219 | -105.8528
19.4257 | -105.8459
19.4294 | -105.8347
19.4331 | -105.8255
19.4368 | -105.8122
19.4405 | -105.7974
19.4442 | -105.7819
19.4479 | -105.7669
19.4516 | -105.7526
19.4553 | -105.7381
19.4591 | -105.7264
19.4628 | -105.7132
19.4665 | -105.696
19.4702 | -105.6789
19.4739 | -105.6626
19.4776 | -105.6474
19.4813 | -105.6321
19.485 | -105.6182
19.4887 | -105.6024
19.4925 | -105.5838
19.4962 | -105.5653
19.4999 | -105.5491
19.5036 | -105.5348
19.5073 | -105.5191
19.511 | -105.5054
19.5147 | -105.4934
19.5184 | -105.4846
19.5222 | -105.4749
19.5259 | -105.4639
19.5296 | -105.4554
19.5333 | -105.4481
19.537 | -105.4406
19.5407 | -105.4327
19.5444 | -105.4256
19.5481 | -105.4194
19.5518 | -105.4141
19.5556 | -105.4106
19.5593 | -105.4072
19.563 | -105.4054
19.5667 | -105.4037
19.5704 | -105.4025
19.5741 | -105.4011
19.5778 | -105.4
19.5815 | -105.4011
19.5852 | -105.4015
19.589 | -105.4015
19.5927 | -105.4024
19.5964 | -105.4041
19.6001 | -105.4049
19.6038 | -105.4057
19.6075 | -105.4078
19.6112 | -105.4094
19.6149 | -105.4107
19.6186 | -105.4151
19.6224 | -105.4193
19.6261 | -105.422
19.6298 | -105.4236
19.6335 | -105.4274
19.6372 | -105.4356
19.6409 | -105.4437
19.6446 | -105.4523
19.6483 | -105.4607
19.6521 | -105.4704
19.6558 | -105.4797
19.6595 | -105.4894
19.6632 | -105.4993
19.6669 | -105.5076
19.6706 | -105.5165
19.6743 | -105.5253
19.678 | -105.532
19.6817 | -105.5406
19.6855 | -105.5528
19.6892 | -105.565
19.6929 | -105.5765
19.6966 | -105.5852
19.7003 | -105.5943
19.704 | -105.6017
19.7077 | -105.6094
19.7114 | -105.618
19.7151 | -105.6271
19.7189 | -105.6336
19.7226 | -105.6379
19.7263 | -105.6435
19.73 | -105.6488
19.7337 | -105.6533
19.7374 | -105.6558
19.7411 | -105.6606
19.7448 | -105.6676
19.7486 | -105.6735
19.7523 | -105.6794
19.756 | -105.6854
19.7597 | -105.6922
19.7634 | -105.6968
19.7671 | -105.7011
19.7708 | -105.705
19.7745 | -105.7107
19.7782 | -105.7143
19.782 | -105.7182
19.7857 | -105.7218
19.7894 | -105.7237
19.7931 | -105.7272
19.7968 | -105.7277
19.8005 | -105.7272
19.8042 | -105.7285
19.8079 | -105.724
19.8116 | -105.7242
19.8154 | -105.7247
19.8191 | -105.7263
19.8228 | -105.7269
19.8265 | -105.7259
19.8302 | -105.7228
19.8339 | -105.7203
19.8376 | -105.7167
19.8413 | -105.7118
19.845 | -105.7069
19.8488 | -105.7019
19.8525 | -105.6977
19.8562 | -105.6957
19.8599 | -105.6921
19.8636 | -105.6881
19.8673 | -105.6848
19.871 | -105.6823
19.8747 | -105.6786
19.8785 | -105.676
19.8822 | -105.6742
19.8859 | -105.6733
19.8896 | -105.6692
19.8933 | -105.6645
19.897 | -105.6634
19.9007 | -105.6621
19.9044 | -105.6617
19.9081 | -105.6606
19.9119 | -105.662
19.9156 | -105.6601
19.9193 | -105.6564
19.923 | -105.6539
19.9267 | -105.6505
19.9304 | -105.648
19.9341 | -105.6438
19.9378 | -105.6411
19.9415 | -105.6392
19.9453 | -105.6365
19.949 | -105.6349
19.9527 | -105.6329
19.9564 | -105.6312
19.9601 | -105.632
19.9638 | -105.6299
19.9675 | -105.6244
19.9712 | -105.6207
19.9749 | -105.617
19.9787 | -105.6136
19.9824 | -105.6089
19.9861 | -105.605
19.9898 | -105.6014
19.9935 | -105.596
19.9972 | -105.5926
20.0009 | -105.5879
20.0046 | -105.5843
20.0084 | -105.5788
20.0121 | -105.5746
20.0158 | -105.5705
20.0195 | -105.5669
20.0232 | -105.5619
20.0269 | -105.5559
20.0306 | -105.5501
20.0343 | -105.543
20.038 | -105.5355
20.0418 | -105.5301
20.0455 | -105.5277
20.0492 | -105.5247
20.0529 | -105.522
20.0566 | -105.5171
20.0603 | -105.5107
20.064 | -105.5053
20.0677 | -105.5018
20.0714 | -105.4989
20.0752 | -105.4928
20.0789 | -105.487
20.0826 | -105.4855
20.0863 | -105.4811
20.09 | -105.4765
20.0937 | -105.4719
20.0974 | -105.4646
20.1011 | -105.4578
20.1048 | -105.4495
20.1086 | -105.4416
20.1123 | -105.4338
20.116 | -105.4278
20.1197 | -105.4212
20.1234 | -105.4154
20.1271 | -105.4074
20.1308 | -105.3986
20.1345 | -105.392
20.1383 | -105.3836
20.142 | -105.3726
20.1457 | -105.3644
20.1494 | -105.3566
20.1531 | -105.3489
20.1568 | -105.3399
20.1605 | -105.3291
20.1642 | -105.3193
20.1679 | -105.3095
20.1717 | -105.3009
20.1754 | -105.294
20.1791 | -105.287
20.1828 | -105.2791
20.1865 | -105.2718
20.1902 | -105.2644
20.1939 | -105.2542
20.1976 | -105.2441
20.2013 | -105.2377
20.2051 | -105.2314
20.2088 | -105.2233
20.2125 | -105.2146
20.2162 | -105.2076
20.2199 | -105.2019
20.2236 | -105.1939
20.2273 | -105.1835
20.231 | -105.1735
20.2347 | -105.1631
20.2385 | -105.155
20.2422 | -105.1489
20.2459 | -105.1403
20.2496 | -105.1315
20.2533 | -105.125
20.257 | -105.1167
20.2607 | -105.1086
20.2644 | -105.1006
20.2682 | -105.0926
20.2719 | -105.0842
20.2756 | -105.075
20.2793 | -105.0674
20.283 | -105.0586
20.2867 | -105.053
20.2904 | -105.0489
20.2941 | -105.0443
20.2978 | -105.0384
20.3016 | -105.0332
20.3053 | -105.0276
20.309 | -105.0232
20.3127 | -105.0217
20.3164 | -105.0202
20.3201 | -105.0184
20.3238 | -105.0178
20.3275 | -105.0143
20.3312 | -105.01
20.335 | -105.0104
20.3387 | -105.0111
20.3424 | -105.0123
20.3461 | -105.0153
20.3498 | -105.0173
20.3535 | -105.0176
20.3572 | -105.0183
20.3609 | -105.0197
20.3646 | -105.0221
20.3684 | -105.0233
20.3721 | -105.0233
20.3758 | -105.0259
20.3795 | -105.0289
20.3832 | -105.0314
20.3869 | -105.0344
20.3906 | -105.0374
20.3943 | -105.0413
20.3981 | -105.0452
20.4018 | -105.0482
20.4055 | -105.0489
20.4092 | -105.0506
20.4129 | -105.0537
20.4166 | -105.0573
20.4203 | -105.0604
20.424 | -105.0629
20.4277 | -105.0662
20.4315 | -105.0725
20.4352 | -105.0779
20.4389 | -105.0833
20.4426 | -105.0876
20.4463 | -105.0947
20.45 | -105.1035
20.4537 | -105.1108
20.4574 | -105.1173
20.4611 | -105.1234
20.4649 | -105.1302
20.4686 | -105.1378
20.4723 | -105.1465
20.476 | -105.1545
20.4797 | -105.1637
20.4834 | -105.177
20.4871 | -105.1892
20.4908 | -105.1993
20.4945 | -105.2082
20.4983 | -105.2194
20.502 | -105.2308
20.5057 | -105.24
20.5094 | -105.2501
20.5131 | -105.2623
20.5168 | -105.276
20.5205 | -105.2893
20.5242 | -105.3012
20.528 | -105.3117
20.5317 | -105.3215
20.5354 | -105.3332
20.5391 | -105.3456
20.5428 | -105.3542
20.5465 | -105.3611
20.5502 | -105.3656
20.5539 | -105.3721
20.5576 | -105.3787
20.5614 | -105.3842
20.5651 | -105.3897
20.5688 | -105.3936
20.5725 | -105.3955
20.5762 | -105.3949
20.5799 | -105.3928
20.5836 | -105.3918
20.5873 | -105.3907
20.591 | -105.3889
20.5948 | -105.3854
20.5985 | -105.3793
20.6022 | -105.3723
20.6059 | -105.3663
20.6096 | -105.3586
20.6133 | -105.348
20.617 | -105.3364
20.6207 | -105.3248
20.6244 | -105.309
20.6282 | -105.2937
20.6319 | -105.2787
20.6356 | -105.2609
20.6393 | -105.2425
20.643 | -105.2237
20.6467 | -105.2041
20.6504 | -105.1828
20.6541 | -105.1596
20.6579 | -105.1377
20.6616 | -105.1169
20.6653 | -105.0933
20.669 | -105.07
20.6727 | -105.0473
20.6764 | -105.0233
20.6801 | -104.9995
20.6838 | -104.9748
20.6875 | -104.9489
20.6913 | -104.9255
20.695 | -104.9044
20.6987 | -104.8855
20.7024 | -104.8654
20.7061 | -104.8439
20.7098 | -104.8243
20.7135 | -104.8054
20.7172 | -104.7871
20.7209 | -104.7692
20.7247 | -104.7509
20.7284 | -104.7341
20.7321 | -104.7179
20.7358 | -104.7033
20.7395 | -104.6897
20.7432 | -104.6806
20.7469 | -104.6721
20.7506 | -104.6651
20.7543 | -104.6591
20.7581 | -104.655
20.7618 | -104.6524
20.7655 | -104.6482
20.7692 | -104.6463
20.7729 | -104.6446
20.7766 | -104.6417
20.7803 | -104.6421
20.784 | -104.6446
20.7878 | -104.6476
20.7915 | -104.6512
20.7952 | -104.6558
20.7989 | -104.6616
20.8026 | -104.6671
20.8063 | -104.6725
20.81 | -104.6769
20.8137 | -104.6811
20.8174 | -104.6878
20.8212 | -104.6928
20.8249 | -104.6977
20.8286 | -104.7028
20.8323 | -104.7077
20.836 | -104.7154
20.8397 | -104.7245
20.8434 | -104.7335
20.8471 | -104.7388
20.8508 | -104.7422
20.8546 | -104.7459
20.8583 | -104.7497
20.862 | -104.7523
20.8657 | -104.758
20.8694 | -104.7631
20.8731 | -104.7639
20.8768 | -104.7644
20.8805 | -104.7645
20.8843 | -104.7658
20.888 | -104.7662
20.8917 | -104.7685
20.8954 | -104.7711
20.8991 | -104.7705
20.9028 | -104.7705
20.9065 | -104.7736
20.9102 | -104.7767
20.9139 | -104.7809
20.9177 | -104.7839
20.9214 | -104.7876
20.9251 | -104.7923
20.9288 | -104.7968
20.9325 | -104.8031
20.9362 | -104.8113
20.9399 | -104.8207
20.9436 | -104.8307
20.9473 | -104.8412
20.9511 | -104.8507
20.9548 | -104.8606
20.9585 | -104.8708
20.9622 | -104.8807
20.9659 | -104.8922
20.9696 | -104.9028
20.9733 | -104.9162
20.977 | -104.9314
20.9807 | -104.9463
20.9845 | -104.9626
20.9882 | -104.9769
20.9919 | -104.9893
20.9956 | -105.0011
20.9993 | -105.0155
21.003 | -105.032
21.0067 | -105.0478
21.0104 | -105.0627
21.0142 | -105.0765
21.0179 | -105.0904
21.0216 | -105.1016
21.0253 | -105.1119
21.029 | -105.1233
21.0327 | -105.135
21.0364 | -105.1441
21.0401 | -105.1499
21.0438 | -105.1538
21.0476 | -105.1572
21.0513 | -105.162
21.055 | -105.1661
21.0587 | -105.1682
21.0624 | -105.1728
21.0661 | -105.175
21.0698 | -105.176
21.0735 | -105.1789
21.0772 | -105.1795
21.081 | -105.1788
21.0847 | -105.1779
21.0884 | -105.1754
21.0921 | -105.1725
21.0958 | -105.169
21.0995 | -105.1644
21.1032 | -105.1579
21.1069 | -105.1526
21.1106 | -105.1482
21.1144 | -105.1424
21.1181 | -105.1375
21.1218 | -105.1294
21.1255 | -105.1217
21.1292 | -105.1158
21.1329 | -105.1101
21.1366 | -105.104
21.1403 | -105.0994
21.1441 | -105.0946
21.1478 | -105.0886
21.1515 | -105.0827
21.1552 | -105.0772
21.1589 | -105.0708
21.1626 | -105.0651
21.1663 | -105.0632
21.17 | -105.0629
21.1737 | -105.0608
21.1775 | -105.059
21.1812 | -105.0545
21.1849 | -105.0498
21.1886 | -105.0448
21.1923 | -105.0408
21.196 | -105.0386
21.1997 | -105.0376
21.2034 | -105.0337
21.2071 | -105.0291
21.2109 | -105.0235
21.2146 | -105.0185
21.2183 | -105.0119
21.222 | -105.0046
21.2257 | -104.9989
21.2294 | -104.9927
21.2331 | -104.9852
21.2368 | -104.9742
21.2405 | -104.9623
21.2443 | -104.9542
21.248 | -104.9464
21.2517 | -104.9391
21.2554 | -104.9296
21.2591 | -104.9196
21.2628 | -104.9117
21.2665 | -104.9021
21.2702 | -104.8933
21.2739 | -104.8831
21.2777 | -104.8735
21.2814 | -104.8649
21.2851 | -104.8539
21.2888 | -104.844
21.2925 | -104.8365
21.2962 | -104.8306
21.2999 | -104.8246
21.3036 | -104.8174
21.3074 | -104.8112
21.3111 | -104.807
21.3148 | -104.8041
21.3185 | -104.8025
21.3222 | -104.8027
21.3259 | -104.8035
21.3296 | -104.8023
21.3333 | -104.8029
21.337 | -104.8051
21.3408 | -104.8089
21.3445 | -104.8117
21.3482 | -104.8176
21.3519 | -104.8249
21.3556 | -104.8314
21.3593 | -104.8399
21.363 | -104.8477
21.3667 | -104.8585
21.3704 | -104.8697
21.3742 | -104.8794
21.3779 | -104.8915
21.3816 | -104.9064
21.3853 | -104.9227
21.389 | -104.9392
21.3927 | -104.9576
21.3964 | -104.9764
21.4001 | -104.9943
21.4039 | -105.0124
21.4076 | -105.0294
21.4113 | -105.0473
21.415 | -105.0621
21.4187 | -105.0782
21.4224 | -105.0953
21.4261 | -105.1149
21.4298 | -105.1325
21.4335 | -105.1474
21.4373 | -105.1601
21.441 | -105.1713
21.4447 | -105.1821
21.4484 | -105.1942
21.4521 | -105.2052
21.4558 | -105.216
21.4595 | -105.2248
21.4632 | -105.2315
21.4669 | -105.2373
21.4707 | -105.2407
21.4744 | -105.2441
21.4781 | -105.2468
21.4818 | -105.2484
21.4855 | -105.2508
21.4892 | -105.2512
21.4929 | -105.2511
21.4966 | -105.2502
21.5003 | -105.2462
21.5041 | -105.2417
21.5078 | -105.2396
21.5115 | -105.2365
21.5152 | -105.2322
21.5189 | -105.2267
21.5226 | -105.2229
21.5263 | -105.2189
21.53 | -105.2132
21.5338 | -105.2092
21.5375 | -105.2065
21.5412 | -105.205
21.5449 | -105.2027
21.5486 | -105.1975
21.5523 | -105.1938
21.556 | -105.1914
21.5597 | -105.1895
21.5634 | -105.1891
21.5672 | -105.1884
21.5709 | -105.1869
21.5746 | -105.1866
21.5783 | -105.1878
21.582 | -105.1872
21.5857 | -105.1857
21.5894 | -105.1875
21.5931 | -105.1889
21.5968 | -105.1903
21.6006 | -105.1906
21.6043 | -105.1889
21.608 | -105.1883
21.6117 | -105.1887
21.6154 | -105.1896
21.6191 | -105.1886
21.6228 | -105.185
21.6265 | -105.1836
21.6302 | -105.1833
21.634 | -105.1833
21.6377 | -105.1815
21.6414 | -105.1783
21.6451 | -105.174
21.6488 | -105.1724
21.6525 | -105.1678
21.6562 | -105.1609
21.6599 | -105.1524
21.6637 | -105.1423
21.6674 | -105.1336
21.6711 | -105.1246
21.6748 | -105.1171
21.6785 | -105.1089
21.6822 | -105.0973
21.6859 | -105.0841
21.6896 | -105.0735
21.6933 | -105.0637
21.6971 | -105.051
21.7008 | -105.0404
21.7045 | -105.0315
21.7082 | -105.0223
21.7119 | -105.0145
21.7156 | -105.0051
21.7193 | -104.9959
21.723 | -104.9879
21.7267 | -104.979
21.7305 | -104.9709
21.7342 | -104.9626
21.7379 | -104.955
21.7416 | -104.9484
21.7453 | -104.9428
21.749 | -104.9378
21.7527 | -104.9313
21.7564 | -104.9265
21.7601 | -104.9234
21.7639 | -104.9171
21.7676 | -104.9142
21.7713 | -104.9148
21.775 | -104.9128
21.7787 | -104.9108
21.7824 | -104.9095
21.7861 | -104.9087
21.7898 | -104.9069
21.7936 | -104.9059
21.7973 | -104.9078
21.801 | -104.9099
21.8047 | -104.9107
21.8084 | -104.9126
21.8121 | -104.9147
21.8158 | -104.9164
21.8195 | -104.9177
21.8232 | -104.9205
21.827 | -104.9222
21.8307 | -104.9236
21.8344 | -104.9256
21.8381 | -104.9264
21.8418 | -104.9253
21.8455 | -104.9248
21.8492 | -104.9251
21.8529 | -104.9258
21.8566 | -104.927
21.8604 | -104.9281
21.8641 | -104.9279
21.8678 | -104.9272
21.8715 | -104.9259
21.8752 | -104.9246
21.8789 | -104.924
21.8826 | -104.922
21.8863 | -104.9194
21.89 | -104.9177
21.8938 | -104.9174
21.8975 | -104.9149
21.9012 | -104.9116
21.9049 | -104.9061
21.9086 | -104.9014
21.9123 | -104.8996
21.916 | -104.8962
21.9197 | -104.8933
21.9235 | -104.8899
21.9272 | -104.8851
21.9309 | -104.8828
21.9346 | -104.8803
21.9383 | -104.8751
21.942 | -104.8691
21.9457 | -104.8661
21.9494 | -104.8639
21.9531 | -104.8654
21.9569 | -104.8638
21.9606 | -104.8625
21.9643 | -104.8609
21.968 | -104.86
21.9717 | -104.8607
21.9754 | -104.8639
21.9791 | -104.8654
21.9828 | -104.8667
21.9865 | -104.8665
21.9903 | -104.8694
21.994 | -104.8733
21.9977 | -104.8773
22.0014 | -104.8774
22.0051 | -104.8776
22.0088 | -104.8787
22.0125 | -104.8787
22.0162 | -104.8782
22.0199 | -104.8799
22.0237 | -104.8805
22.0274 | -104.881
22.0311 | -104.88
22.0348 | -104.8804
22.0385 | -104.8805
22.0422 | -104.88
22.0459 | -104.8809
22.0496 | -104.8823
22.0534 | -104.8827
22.0571 | -104.884
22.0608 | -104.8867
22.0645 | -104.8892
22.0682 | -104.8886
22.0719 | -104.8855
22.0756 | -104.8849
22.0793 | -104.8816
22.083 | -104.8782
22.0868 | -104.8759
22.0905 | -104.8709
22.0942 | -104.8668
22.0979 | -104.8634
22.1016 | -104.861
22.1053 | -104.8568
22.109 | -104.8526
22.1127 | -104.8492
22.1164 | -104.8451
22.1202 | -104.8404
22.1239 | -104.8358
22.1276 | -104.831
22.1313 | -104.8266
22.135 | -104.8227
22.1387 | -104.8179
22.1424 | -104.8132
22.1461 | -104.8103
22.1498 | -104.8082
22.1536 | -104.8061
22.1573 | -104.8031
22.161 | -104.8015
22.1647 | -104.8012
22.1684 | -104.7999
22.1721 | -104.7992
22.1758 | -104.7991
22.1795 | -104.8011
22.1833 | -104.8034
22.187 | -104.805
22.1907 | -104.8074
22.1944 | -104.8098
22.1981 | -104.8124
22.2018 | -104.8151
22.2055 | -104.8187
22.2092 | -104.8243
22.2129 | -104.8305
22.2167 | -104.8377
22.2204 | -104.8449
22.2241 | -104.8499
22.2278 | -104.8546
22.2315 | -104.8573
22.2352 | -104.8594
22.2389 | -104.8666
22.2426 | -104.8722
22.2463 | -104.8757
22.2501 | -104.8818
22.2538 | -104.8903
22.2575 | -104.8965
22.2612 | -104.9015
22.2649 | -104.9067
22.2686 | -104.9108
22.2723 | -104.9171
22.276 | -104.9231
22.2797 | -104.9279
22.2835 | -104.9332
22.2872 | -104.9407
22.2909 | -104.9467
22.2946 | -104.9538
22.2983 | -104.9594
22.302 | -104.9671
22.3057 | -104.9729
22.3094 | -104.9792
22.3132 | -104.9872
22.3169 | -104.9961
22.3206 | -105.0061
22.3243 | -105.0171
22.328 | -105.0276
22.3317 | -105.0395
22.3354 | -105.0484
22.3391 | -105.0568
22.3428 | -105.066
22.3466 | -105.0759
22.3503 | -105.0847
22.354 | -105.0939
22.3577 | -105.1029
22.3614 | -105.1089
22.3651 | -105.1153
22.3688 | -105.1224
22.3725 | -105.1289
22.3762 | -105.1356
22.38 | -105.144
22.3837 | -105.1528
22.3874 | -105.1629
22.3911 | -105.1718
22.3948 | -105.18
22.3985 | -105.1876
22.4022 | -105.1956
22.4059 | -105.2042
22.4096 | -105.2075
22.4134 | -105.2062
22.4171 | -105.2052
22.4208 | -105.2088
22.4245 | -105.2127
22.4282 | -105.2135
22.4319 | -105.2095
22.4356 | -105.1991
22.4393 | -105.1852
22.4431 | -105.1761
22.4468 | -105.1676
22.4505 | -105.1597
22.4542 | -105.153
22.4579 | -105.1429
22.4616 | -105.1275
22.4653 | -105.1112
22.469 | -105.0901
22.4727 | -105.0632
22.4765 | -105.0373
22.4802 | -105.0101
22.4839 | -104.9784
22.4876 | -104.9303
22.4913 | -104.898
22.495 | -104.8646
22.4987 | -104.8295
22.5024 | -104.7888
22.5061 | -104.7408
22.5099 | -104.6918
22.5136 | -104.6441
22.5173 | -104.5979
22.521 | -104.5507
22.5247 | -104.5027
22.5284 | -104.4541
22.5321 | -104.4067
22.5358 | -104.3585
22.5395 | -104.3089
22.5433 | -104.2612
22.547 | -104.2149
22.5507 | -104.1684
22.5544 | -104.1231
22.5581 | -104.077
22.5618 | -104.0324
22.5655 | -103.99
22.5692 | -103.9484
22.573 | -103.909
22.5767 | -103.8701
22.5804 | -103.8342
22.5841 | -103.8013
22.5878 | -103.7707
22.5915 | -103.7416
22.5952 | -103.715
22.5989 | -103.6913
22.6026 | -103.6697
22.6064 | -103.6506
22.6101 | -103.6346
22.6138 | -103.6213
22.6175 | -103.6118
22.6212 | -103.6052
22.6249 | -103.6019
22.6286 | -103.6008
22.6323 | -103.6023
22.636 | -103.6072
22.6398 | -103.6139
22.6435 | -103.6244
22.6472 | -103.6376
22.6509 | -103.6513
22.6546 | -103.6686
22.6583 | -103.6894
22.662 | -103.7124
22.6657 | -103.7379
22.6695 | -103.7666
22.6732 | -103.7965
22.6769 | -103.8279
22.6806 | -103.862
22.6843 | -103.897
22.688 | -103.9335
22.6917 | -103.9716
22.6954 | -104.011
22.6991 | -104.0508
22.7029 | -104.091
22.7066 | -104.1325
22.7103 | -104.1748
22.714 | -104.2166
22.7177 | -104.2593
22.7214 | -104.3024
22.7251 | -104.345
22.7288 | -104.3911
22.7325 | -104.4386
22.7363 | -104.4824
22.74 | -104.5189
22.7437 | -104.5535
22.7474 | -104.5878
22.7511 | -104.6249
22.7548 | -104.6645
22.7585 | -104.7042
22.7622 | -104.7352
22.7659 | -104.7624
22.7697 | -104.7856
22.7734 | -104.8062
22.7771 | -104.8261
22.7808 | -104.8466
22.7845 | -104.8648
22.7882 | -104.884
22.7919 | -104.9045
22.7956 | -104.9198
22.7994 | -104.9317
22.8031 | -104.944
22.8068 | -104.9523
22.8105 | -104.9587
22.8142 | -104.9641
22.8179 | -104.969
22.8216 | -104.9714
22.8253 | -104.9755
22.829 | -104.9839
22.8328 | -104.9875
22.8365 | -104.9863
22.8402 | -104.9824
22.8439 | -104.9769
22.8476 | -104.97
22.8513 | -104.962
22.855 | -104.9522
22.8587 | -104.9433
22.8624 | -104.9353
22.8662 | -104.9271
22.8699 | -104.9186
22.8736 | -104.9128
22.8773 | -104.9055
22.881 | -104.8977
22.8847 | -104.8889
22.8884 | -104.8779
22.8921 | -104.8669
22.8958 | -104.8561
22.8996 | -104.8474
22.9033 | -104.8397
22.907 | -104.831
22.9107 | -104.8226
22.9144 | -104.8135
22.9181 | -104.8083
22.9218 | -104.8033
22.9255 | -104.7974
22.9293 | -104.7908
22.933 | -104.7839
22.9367 | -104.7772
22.9404 | -104.7704
22.9441 | -104.7655
22.9478 | -104.761
22.9515 | -104.7576
22.9552 | -104.7538
22.9589 | -104.7499
22.9627 | -104.7486
22.9664 | -104.7461
22.9701 | -104.7455
22.9738 | -104.7461
22.9775 | -104.7474
22.9812 | -104.7485
22.9849 | -104.7495
22.9886 | -104.7504
22.9923 | -104.7485
22.9961 | -104.7506
22.9998 | -104.7539
23.0035 | -104.7579
23.0072 | -104.7633
23.0109 | -104.7669
23.0146 | -104.7691
23.0183 | -104.7722
23.022 | -104.7805
23.0257 | -104.7867
23.0295 | -104.7947
23.0332 | -104.8012
23.0369 | -104.8071
23.0406 | -104.8154
23.0443 | -104.8238
23.048 | -104.8309
23.0517 | -104.8393
23.0554 | -104.8472
23.0592 | -104.8552
23.0629 | -104.8651
23.0666 | -104.8772
23.0703 | -104.8867
23.074 | -104.8965
23.0777 | -104.906
23.0814 | -104.9146
23.0851 | -104.9209
23.0888 | -104.9275
23.0926 | -104.9347
23.0963 | -104.9417
23.1 | -104.9478
23.1037 | -104.9554
23.1074 | -104.9624
23.1111 | -104.9689
23.1148 | -104.976
23.1185 | -104.9832
23.1222 | -104.9879
23.126 | -104.9912
23.1297 | -104.9942
23.1334 | -104.9975
23.1371 | -104.9991
23.1408 | -105
23.1445 | -104.9994
23.1482 | -104.9994
23.1519 | -105.0015
23.1556 | -104.9995
23.1594 | -104.9976
23.1631 | -104.9973
23.1668 | -104.9951
23.1705 | -104.9924
23.1742 | -104.9887
23.1779 | -104.9829
23.1816 | -104.9764
23.1853 | -104.969
23.1891 | -104.9605
23.1928 | -104.952
23.1965 | -104.9449
23.2002 | -104.9357
23.2039 | -104.9257
23.2076 | -104.9166
23.2113 | -104.9088
23.215 | -104.9019
23.2187 | -104.8931
23.2225 | -104.8826
23.2262 | -104.8715
23.2299 | -104.8607
23.2336 | -104.8496
23.2373 | -104.8376
23.241 | -104.827
23.2447 | -104.8153
23.2484 | -104.8045
23.2521 | -104.793
23.2559 | -104.7811
23.2596 | -104.7717
23.2633 | -104.7629
23.267 | -104.7539
23.2707 | -104.7443
23.2744 | -104.7353
23.2781 | -104.7266
23.2818 | -104.7187
23.2855 | -104.7125
23.2893 | -104.706
23.293 | -104.7003
23.2967 | -104.6947
23.3004 | -104.6899
23.3041 | -104.6849
23.3078 | -104.6802
23.3115 | -104.6759
23.3152 | -104.6739
23.319 | -104.6718
23.3227 | -104.6673
23.3264 | -104.6641
23.3301 | -104.6612
23.3338 | -104.6563
23.3375 | -104.6548
23.3412 | -104.6545
23.3449 | -104.6539
23.3486 | -104.6542
23.3524 | -104.6549
23.3561 | -104.6538
23.3598 | -104.6526
23.3635 | -104.6538
23.3672 | -104.6556
23.3709 | -104.6561
23.3746 | -104.6569
23.3783 | -104.6589
23.382 | -104.6627
23.3858 | -104.6661
23.3895 | -104.6692
23.3932 | -104.6709
23.3969 | -104.6731
23.4006 | -104.6772
23.4043 | -104.682
23.408 | -104.6855
23.4117 | -104.6902
23.4154 | -104.6957
23.4192 | -104.7004
23.4229 | -104.704
23.4266 | -104.7078
23.4303 | -104.712
23.434 | -104.7154
23.4377 | -104.719
23.4414 | -104.7217
23.4451 | -104.7236
23.4489 | -104.7247
23.4526 | -104.7281
23.4563 | -104.7324
23.46 | -104.7378
23.4637 | -104.7441
23.4674 | -104.7495
23.4711 | -104.7542
23.4748 | -104.7605
23.4785 | -104.7662
23.4823 | -104.7707
23.486 | -104.7777
23.4897 | -104.7831
23.4934 | -104.7875
23.4971 | -104.7907
23.5008 | -104.7924
23.5045 | -104.7955
23.5082 | -104.8009
23.5119 | -104.8051
23.5157 | -104.8086
23.5194 | -104.8137
23.5231 | -104.8187
23.5268 | -104.8216
23.5305 | -104.825
23.5342 | -104.8289
23.5379 | -104.831
23.5416 | -104.8297
23.5453 | -104.8294
23.5491 | -104.8318
23.5528 | -104.8331
23.5565 | -104.8325
23.5602 | -104.8345
23.5639 | -104.8359
23.5676 | -104.8373
23.5713 | -104.8403
23.575 | -104.8419
23.5788 | -104.8422
23.5825 | -104.8432
23.5862 | -104.8451
23.5899 | -104.8478
23.5936 | -104.8491
23.5973 | -104.8488
23.601 | -104.8503
23.6047 | -104.8518
23.6084 | -104.8535
23.6122 | -104.856
23.6159 | -104.8583
23.6196 | -104.8597
23.6233 | -104.8619
23.627 | -104.8654
23.6307 | -104.867
23.6344 | -104.8683
23.6381 | -104.872
23.6418 | -104.8757
23.6456 | -104.8807
23.6493 | -104.884
23.653 | -104.8901
23.6567 | -104.8967
23.6604 | -104.9005
23.6641 | -104.9053
23.6678 | -104.9135
23.6715 | -104.9197
23.6752 | -104.9251
23.679 | -104.9284
23.6827 | -104.9313
23.6864 | -104.9362
23.6901 | -104.9423
23.6938 | -104.9478
23.6975 | -104.9527
23.7012 | -104.9574
23.7049 | -104.9633
23.7087 | -104.9682
23.7124 | -104.974
23.7161 | -104.979
23.7198 | -104.9836
23.7235 | -104.9895
23.7272 | -104.9956
23.7309 | -105.0017
23.7346 | -105.0078
23.7383 | -105.0137
23.7421 | -105.0192
23.7458 | -105.0227
23.7495 | -105.026
23.7532 | -105.0292
23.7569 | -105.0341
23.7606 | -105.038
23.7643 | -105.0421
23.768 | -105.0449
23.7717 | -105.0471
23.7755 | -105.0503
23.7792 | -105.0539
23.7829 | -105.0577
23.7866 | -105.0577
23.7903 | -105.0583
23.794 | -105.0574
23.7977 | -105.0572
23.8014 | -105.0589
23.8052 | -105.0597
23.8089 | -105.0588
23.8126 | -105.0586
23.8163 | -105.0573
23.82 | -105.0566
23.8237 | -105.0575
23.8274 | -105.0564
23.8311 | -105.057
23.8348 | -105.0589
23.8386 | -105.0592
23.8423 | -105.0572
23.846 | -105.0572
23.8497 | -105.0579
23.8534 | -105.0567
23.8571 | -105.0559
23.8608 | -105.0561
23.8645 | -105.0567
23.8682 | -105.0553
23.872 | -105.0533
23.8757 | -105.0531
23.8794 | -105.0521
23.8831 | -105.0519
23.8868 | -105.0536
23.8905 | -105.0538
23.8942 | -105.0529
23.8979 | -105.0516
23.9016 | -105.0491
23.9054 | -105.0478
23.9091 | -105.0466
23.9128 | -105.044
23.9165 | -105.0398
23.9202 | -105.0356
23.9239 | -105.0326
23.9276 | -105.0293
23.9313 | -105.0238
23.9351 | -105.0212
23.9388 | -105.0179
23.9425 | -105.0144
23.9462 | -105.0099
23.9499 | -105.0045
23.9536 | -104.998
23.9573 | -104.9935
23.961 | -104.9884
23.9647 | -104.982
23.9685 | -104.9788
23.9722 | -104.9767
23.9759 | -104.9733
23.9796 | -104.97
23.9833 | -104.9641
23.987 | -104.9595
23.9907 | -104.9549
23.9944 | -104.951
23.9981 | -104.9478
24.0019 | -104.9439
24.0056 | -104.9402
24.0093 | -104.9369
24.013 | -104.9344
24.0167 | -104.9318
24.0204 | -104.9279
24.0241 | -104.9248
24.0278 | -104.9222
24.0315 | -104.922
24.0353 | -104.9224
24.039 | -104.9235
24.0427 | -104.9229
24.0464 | -104.9213
24.0501 | -104.9201
24.0538 | -104.9207
24.0575 | -104.9222
24.0612 | -104.924
24.0649 | -104.9257
24.0687 | -104.926
24.0724 | -104.927
24.0761 | -104.9299
24.0798 | -104.9329
24.0835 | -104.9364
24.0872 | -104.9388
24.0909 | -104.9413
24.0946 | -104.9455
24.0984 | -104.9466
24.1021 | -104.9475
24.1058 | -104.9499
24.1095 | -104.9528
24.1132 | -104.956
24.1169 | -104.9593
24.1206 | -104.9627
24.1243 | -104.9648
24.128 | -104.9665
24.1318 | -104.9687
24.1355 | -104.9721
24.1392 | -104.9735
24.1429 | -104.973
24.1466 | -104.9743
24.1503 | -104.9757
24.154 | -104.9779
24.1577 | -104.9815
24.1614 | -104.9825
24.1652 | -104.9817
24.1689 | -104.9826
24.1726 | -104.9833
24.1763 | -104.9827
24.18 | -104.9831
24.1837 | -104.9858
24.1874 | -104.9879
24.1911 | -104.9879
24.1948 | -104.9864
24.1986 | -104.9822
24.2023 | -104.9796
24.206 | -104.9778
24.2097 | -104.9773
24.2134 | -104.9764
24.2171 | -104.9732
24.2208 | -104.9704
24.2245 | -104.9646
24.2283 | -104.9573
24.232 | -104.9523
24.2357 | -104.9473
24.2394 | -104.943
24.2431 | -104.9374
24.2468 | -104.9297
24.2505 | -104.9217
24.2542 | -104.9157
24.2579 | -104.9093
24.2617 | -104.891
24.2654 | -104.8608
24.2691 | -104.8257
24.2728 | -104.7782
24.2765 | -104.7189
24.2802 | -104.6491
24.2839 | -104.5875
24.2876 | -104.5336
24.2913 | -104.4733
24.2951 | -104.416
24.2988 | -104.3527
24.3025 | -104.2909
24.3062 | -104.2303
24.3099 | -104.1674
24.3136 | -104.102
24.3173 | -104.033
24.321 | -103.9608
24.3248 | -103.8891
24.3285 | -103.8129
24.3322 | -103.7327
24.3359 | -103.6501
24.3396 | -103.5653
24.3433 | -103.4786
24.347 | -103.3917
24.3507 | -103.3036
24.3544 | -103.2151
24.3582 | -103.127
24.3619 | -103.0391
24.3656 | -102.9513
24.3693 | -102.8635
24.373 | -102.7754
24.3767 | -102.6877
24.3804 | -102.6022
24.3841 | -102.5205
24.3878 | -102.4419
24.3916 | -102.3658
24.3953 | -102.2914
24.399 | -102.2208
24.4027 | -102.1537
24.4064 | -102.091
24.4101 | -102.0329
24.4138 | -101.9788
24.4175 | -101.9282
24.4212 | -101.883
24.425 | -101.8432
24.4287 | -101.8089
24.4324 | -101.7798
24.4361 | -101.7556
24.4398 | -101.7368
24.4435 | -101.7229
24.4472 | -101.7148
24.4509 | -101.7123
24.4547 | -101.715
24.4584 | -101.7227
24.4621 | -101.736
24.4658 | -101.7553
24.4695 | -101.7799
24.4732 | -101.8092
24.4769 | -101.8433
24.4806 | -101.8824
24.4843 | -101.9255
24.4881 | -101.9737
24.4918 | -102.0256
24.4955 | -102.0817
24.4992 | -102.1418
24.5029 | -102.2058
24.5066 | -102.2745
24.5103 | -102.3463
24.514 | -102.4202
24.5177 | -102.4955
24.5215 | -102.5735
24.5252 | -102.6526
24.5289 | -102.732
24.5326 | -102.8123
24.5363 | -102.8951
24.54 | -102.9792
24.5437 | -103.0632
24.5474 | -103.1451
24.5511 | -103.2229
24.5549 | -103.2952
24.5586 | -103.3568
24.5623 | -103.4065
24.566 | -103.4394
24.5697 | -103.4613
24.5734 | -103.4649
24.5771 | -103.438
24.5808 | -103.3955
24.5846 | -103.3561
24.5883 | -103.3253
24.592 | -103.2928
24.5957 | -103.2572
24.5994 | -103.217
24.6031 | -103.1688
24.6068 | -103.1151
24.6105 | -103.0534
24.6142 | -102.9807
24.618 | -102.9026
24.6217 | -102.8172
24.6254 | -102.727
24.6291 | -102.5408
24.6328 | -102.4351
24.6365 | -102.3194
24.6402 | -102.2009
24.6439 | -102.0765
24.6476 | -101.9501
24.6514 | -101.8184
24.6551 | -101.6848
24.6588 | -101.5521
24.6625 | -101.4201
24.6662 | -101.2897
24.6699 | -101.1625
24.6736 | -101.0367
24.6773 | -100.9149
24.681 | -100.7974
24.6848 | -100.6832
24.6885 | -100.5732
24.6922 | -100.4674
24.6959 | -100.3655
24.6996 | -100.2692
24.7033 | -100.1786
24.707 | -100.0939
24.7107 | -100.0148
24.7145 | -99.9416
24.7182 | -99.8747
24.7219 | -99.8145
24.7256 | -99.7609
24.7293 | -99.7136
24.733 | -99.673
24.7367 | -99.6391
24.7404 | -99.6117
24.7441 | -99.5917
24.7479 | -99.5788
24.7516 | -99.5728
24.7553 | -99.5738
24.759 | -99.5818
24.7627 | -99.5968
24.7664 | -99.6186
24.7701 | -99.6468
24.7738 | -99.6818
24.7775 | -99.7236
24.7813 | -99.7718
24.785 | -99.8264
24.7887 | -99.8866
24.7924 | -99.9543
24.7961 | -100.028
24.7998 | -100.1062
24.8035 | -100.1901
24.8072 | -100.2795
24.8109 | -100.3735
24.8147 | -100.4722
24.8184 | -100.5755
24.8221 | -100.6836
24.8258 | -100.7958
24.8295 | -100.9109
24.8332 | -101.0294
24.8369 | -101.1503
24.8406 | -101.2734
24.8444 | -101.399
24.8481 | -101.5255
24.8518 | -101.6526
24.8555 | -101.7797
24.8592 | -101.9077
24.8629 | -102.0359
24.8666 | -102.164
24.8703 | -102.2903
24.874 | -102.4152
24.8778 | -102.5383
24.8815 | -102.6592
24.8852 | -102.7771
24.8889 | -102.8917
24.8926 | -103.0029
24.8963 | -103.1114
24.9 | -103.2175
24.9037 | -103.3236
24.9074 | -103.4326
24.9112 | -103.5445
24.9149 | -103.6587
24.9186 | -103.7679
24.9223 | -103.8796
24.926 | -104.0019
24.9297 | -104.1179
24.9334 | -104.207
24.9371 | -104.2668
24.9408 | -104.3119
24.9446 | -104.3447
24.9483 | -104.3693
24.952 | -104.3916
24.9557 | -104.4163
24.9594 | -104.4433
24.9631 | -104.4693
24.9668 | -104.4955
24.9705 | -104.5198
24.9743 | -104.5381
24.978 | -104.5541
24.9817 | -104.568
24.9854 | -104.5779
24.9891 | -104.5847
24.9928 | -104.5919
24.9965 | -104.5971
25.0002 | -104.6026
25.0039 | -104.6106
25.0077 | -104.6177
25.0114 | -104.623
25.0151 | -104.6287
25.0188 | -104.6354
25.0225 | -104.6406
25.0262 | -104.6464
25.0299 | -104.6519
25.0336 | -104.6554
25.0373 | -104.6595
25.0411 | -104.6654
25.0448 | -104.6716
25.0485 | -104.6757
25.0522 | -104.6787
25.0559 | -104.6808
25.0596 | -104.6812
25.0633 | -104.6831
25.067 | -104.6851
25.0707 | -104.6883
25.0745 | -104.6909
25.0782 | -104.6948
25.0819 | -104.6957
25.0856 | -104.6966
25.0893 | -104.6977
25.093 | -104.7
25.0967 | -104.7017
25.1004 | -104.7018
25.1042 | -104.7011
25.1079 | -104.7006
25.1116 | -104.6994
25.1153 | -104.6987
25.119 | -104.6983
25.1227 | -104.6971
25.1264 | -104.6939
25.1301 | -104.6931
25.1338 | -104.6922
25.1376 | -104.691
25.1413 | -104.6902
25.145 | -104.6873
25.1487 | -104.6841
25.1524 | -104.683
25.1561 | -104.68
25.1598 | -104.6765
25.1635 | -104.6751
25.1672 | -104.6747
25.171 | -104.6728
25.1747 | -104.6698
25.1784 | -104.667
25.1821 | -104.6634
25.1858 | -104.6595
25.1895 | -104.6556
25.1932 | -104.6535
25.1969 | -104.6507
25.2006 | -104.6477
25.2044 | -104.6441
25.2081 | -104.64
25.2118 | -104.636
25.2155 | -104.6317
25.2192 | -104.6281
25.2229 | -104.6244
25.2266 | -104.6218
25.2303 | -104.6188
25.2341 | -104.6154
25.2378 | -104.6123
25.2415 | -104.6088
25.2452 | -104.605
25.2489 | -104.5998
25.2526 | -104.5939
25.2563 | -104.5898
25.26 | -104.5849
25.2637 | -104.5791
25.2675 | -104.5735
25.2712 | -104.5676
25.2749 | -104.5636
25.2786 | -104.5582
25.2823 | -104.5537
25.286 | -104.5484
25.2897 | -104.5435
25.2934 | -104.54
25.2971 | -104.5361
25.3009 | -104.5333
25.3046 | -104.5294
25.3083 | -104.5251
25.312 | -104.5209
25.3157 | -104.5178
25.3194 | -104.513
25.3231 | -104.5067
25.3268 | -104.5014
25.3305 | -104.4977
25.3343 | -104.4931
25.338 | -104.4866
25.3417 | -104.4815
25.3454 | -104.4777
25.3491 | -104.4729
25.3528 | -104.4685
25.3565 | -104.4646
25.3602 | -104.4584
25.364 | -104.4535
25.3677 | -104.4485
25.3714 | -104.4443
25.3751 | -104.4401
25.3788 | -104.4363
25.3825 | -104.4318
25.3862 | -104.4288
25.3899 | -104.4268
25.3936 | -104.4241
25.3974 | -104.4207
25.4011 | -104.4169
25.4048 | -104.4133
25.4085 | -104.4113
25.4122 | -104.4097
25.4159 | -104.4095
25.4196 | -104.4079
25.4233 | -104.406
25.427 | -104.4037
25.4308 | -104.4022
25.4345 | -104.4006
25.4382 | -104.3999
25.4419 | -104.3983
25.4456 | -104.3974
25.4493 | -104.3958
25.453 | -104.3937
25.4567 | -104.3932
25.4605 | -104.3944
25.4642 | -104.3942
25.4679 | -104.3955
25.4716 | -104.3958
25.4753 | -104.3969
25.479 | -104.3983
25.4827 | -104.3997
25.4864 | -104.3992
25.4901 | -104.4003
25.4939 | -104.4018
25.4976 | -104.4038
25.5013 | -104.4038
25.505 | -104.4033
25.5087 | -104.4028
25.5124 | -104.4049
25.5161 | -104.4068
25.5198 | -104.4055
25.5235 | -104.4034
25.5273 | -104.4016
25.531 | -104.4017
25.5347 | -104.4011
25.5384 | -104.4008
25.5421 | -104.401
25.5458 | -104.3994
25.5495 | -104.3978
25.5532 | -104.3967
25.5569 | -104.3963
25.5607 | -104.3935
25.5644 | -104.3913
25.5681 | -104.3888
25.5718 | -104.3879
25.5755 | -104.3859
25.5792 | -104.3825
25.5829 | -104.38
25.5866 | -104.3755
25.5904 | -104.3696
25.5941 | -104.3646
25.5978 | -104.3595
25.6015 | -104.3542
25.6052 | -104.3491
25.6089 | -104.3464
25.6126 | -104.3438
25.6163 | -104.3386
25.62 | -104.334
25.6238 | -104.3292
25.6275 | -104.3237
25.6312 | -104.3198
25.6349 | -104.3144
25.6386 | -104.3111
25.6423 | -104.3055
25.646 | -104.3002
25.6497 | -104.2956
25.6534 | -104.2886
25.6572 | -104.2827
25.6609 | -104.2765
25.6646 | -104.2711
25.6683 | -104.266
25.672 | -104.262
25.6757 | -104.2575
25.6794 | -104.2556
25.6831 | -104.2543
25.6868 | -104.2513
25.6906 | -104.247
25.6943 | -104.2434
25.698 | -104.2403
25.7017 | -104.2372
25.7054 | -104.2331
25.7091 | -104.2287
25.7128 | -104.2267
25.7165 | -104.2247
25.7203 | -104.2211
25.724 | -104.2177
25.7277 | -104.2155
25.7314 | -104.2136
25.7351 | -104.2107
25.7388 | -104.208
25.7425 | -104.2057
25.7462 | -104.2045
25.7499 | -104.2021
25.7537 | -104.1995
25.7574 | -104.1979
25.7611 | -104.1958
25.7648 | -104.1934
25.7685 | -104.1915
25.7722 | -104.1889
25.7759 | -104.1852
25.7796 | -104.1807
25.7833 | -104.1776
25.7871 | -104.1744
25.7908 | -104.1725
25.7945 | -104.1695
25.7982 | -104.1673
25.8019 | -104.1672
25.8056 | -104.1672
25.8093 | -104.1658
25.813 | -104.1643
25.8167 | -104.1635
25.8205 | -104.1616
25.8242 | -104.1567
25.8279 | -104.1538
25.8316 | -104.1511
25.8353 | -104.1483
25.839 | -104.1457
25.8427 | -104.1457
25.8464 | -104.1463
25.8502 | -104.1457
25.8539 | -104.1448
25.8576 | -104.1443
25.8613 | -104.1448
25.865 | -104.1484
25.8687 | -104.153
25.8724 | -104.1567
25.8761 | -104.1588
25.8798 | -104.1609
25.8836 | -104.1653
25.8873 | -104.1699
25.891 | -104.174
25.8947 | -104.1792
25.8984 | -104.187
25.9021 | -104.1945
25.9058 | -104.2009
25.9095 | -104.2094
25.9132 | -104.2184
25.917 | -104.2262
25.9207 | -104.2347
25.9244 | -104.2442
25.9281 | -104.2539
25.9318 | -104.2639
25.9355 | -104.2749
25.9392 | -104.2864
25.9429 | -104.2973
25.9466 | -104.3078
25.9504 | -104.3205
25.9541 | -104.3349
25.9578 | -104.3459
25.9615 | -104.3569
25.9652 | -104.3703
25.9689 | -104.3839
25.9726 | -104.3958
25.9763 | -104.4064
25.9801 | -104.418
25.9838 | -104.4293
25.9875 | -104.4419
25.9912 | -104.4529
25.9949 | -104.4637
25.9986 | -104.475
26.0023 | -104.4844
26.006 | -104.4962
26.0097 | -104.508
26.0135 | -104.5181
26.0172 | -104.5284
26.0209 | -104.5363
26.0246 | -104.5448
26.0283 | -104.5547
26.032 | -104.5628
26.0357 | -104.5699
26.0394 | -104.5769
26.0431 | -104.5838
26.0469 | -104.5873
26.0506 | -104.5897
26.0543 | -104.5921
26.058 | -104.5938
26.0617 | -104.5972
26.0654 | -104.6016
26.0691 | -104.6027
26.0728 | -104.6025
26.0765 | -104.604
26.0803 | -104.6045
26.084 | -104.6044
26.0877 | -104.606
26.0914 | -104.606
26.0951 | -104.604
26.0988 | -104.6016
26.1025 | -104.5994
26.1062 | -104.5964
26.11 | -104.5931
26.1137 | -104.5886
26.1174 | -104.5833
26.1211 | -104.5781
26.1248 | -104.5737
26.1285 | -104.5694
26.1322 | -104.5641
26.1359 | -104.5586
26.1396 | -104.5528
26.1434 | -104.5463
26.1471 | -104.5411
26.1508 | -104.535
26.1545 | -104.5281
26.1582 | -104.5212
26.1619 | -104.5153
26.1656 | -104.5097
26.1693 | -104.5042
26.173 | -104.498
26.1768 | -104.4924
26.1805 | -104.4898
26.1842 | -104.4856
26.1879 | -104.4821
26.1916 | -104.4819
26.1953 | -104.4807
26.199 | -104.4782
26.2027 | -104.4747
26.2064 | -104.4724
26.2102 | -104.4698
26.2139 | -104.466
26.2176 | -104.4632
26.2213 | -104.4614
26.225 | -104.4578
26.2287 | -104.4549
26.2324 | -104.4517
26.2361 | -104.4497
26.2399 | -104.4467
26.2436 | -104.4438
26.2473 | -104.4439
26.251 | -104.4444
26.2547 | -104.4434
26.2584 | -104.4436
26.2621 | -104.4448
26.2658 | -104.4438
26.2695 | -104.4432
26.2733 | -104.4427
26.277 | -104.4435
26.2807 | -104.445
26.2844 | -104.4462
26.2881 | -104.4478
26.2918 | -104.4496
26.2955 | -104.4509
26.2992 | -104.4535
26.3029 | -104.4561
26.3067 | -104.4567
26.3104 | -104.458
26.3141 | -104.4599
26.3178 | -104.4601
26.3215 | -104.4622
26.3252 | -104.4646
26.3289 | -104.4678
26.3326 | -104.4695
26.3363 | -104.4697
26.3401 | -104.4713
26.3438 | -104.4747
26.3475 | -104.4784
26.3512 | -104.4804
26.3549 | -104.4803
26.3586 | -104.4797
26.3623 | -104.4822
26.366 | -104.4837
26.3698 | -104.4843
26.3735 | -104.4864
26.3772 | -104.4882
26.3809 | -104.4897
26.3846 | -104.4906
26.3883 | -104.4939
26.392 | -104.4977
26.3957 | -104.4993
26.3994 | -104.5003
26.4032 | -104.503
26.4069 | -104.5062
26.4106 | -104.509
26.4143 | -104.5106
26.418 | -104.5147
26.4217 | -104.519
26.4254 | -104.5218
26.4291 | -104.5254
26.4328 | -104.5276
26.4366 | -104.529
26.4403 | -104.5329
26.444 | -104.5364
26.4477 | -104.5366
26.4514 | -104.5376
26.4551 | -104.5402
26.4588 | -104.5423
26.4625 | -104.5459
26.4662 | -104.5494
26.47 | -104.5532
26.4737 | -104.5557
26.4774 | -104.5582
26.4811 | -104.5612
26.4848 | -104.564
26.4885 | -104.5672
26.4922 | -104.5706
26.4959 | -104.5738
26.4997 | -104.5769
26.5034 | -104.5815
26.5071 | -104.5838
26.5108 | -104.5851
26.5145 | -104.5875
26.5182 | -104.5885
26.5219 | -104.5881
26.5256 | -104.5896
26.5293 | -104.5925
26.5331 | -104.5956
26.5368 | -104.5977
26.5405 | -104.6002
26.5442 | -104.6016
26.5479 | -104.6041
26.5516 | -104.6067
26.5553 | -104.6114
26.559 | -104.6154
26.5627 | -104.618
26.5665 | -104.6217
26.5702 | -104.6257
26.5739 | -104.6307
26.5776 | -104.6352
26.5813 | -104.6391
26.585 | -104.6452
26.5887 | -104.6504
26.5924 | -104.6548
26.5961 | -104.6624
26.5999 | -104.6685
26.6036 | -104.6739
26.6073 | -104.6816
26.611 | -104.689
26.6147 | -104.6949
26.6184 | -104.7017
26.6221 | -104.7076
26.6258 | -104.7127
26.6296 | -104.7184
26.6333 | -104.7228
26.637 | -104.7276
26.6407 | -104.7312
26.6444 | -104.734
26.6481 | -104.7388
26.6518 | -104.7428
26.6555 | -104.7461
26.6592 | -104.7489
26.663 | -104.7511
26.6667 | -104.7544
26.6704 | -104.7588
26.6741 | -104.7618
26.6778 | -104.7656
26.6815 | -104.7685
26.6852 | -104.7706
26.6889 | -104.7743
26.6926 | -104.779
26.6964 | -104.78
26.7001 | -104.7796
26.7038 | -104.7808
26.7075 | -104.7832
26.7112 | -104.7846
26.7149 | -104.7848
26.7186 | -104.7844
26.7223 | -104.7842
26.7261 | -104.7865
26.7298 | -104.7891
26.7335 | -104.7901
26.7372 | -104.7901
26.7409 | -104.7904
26.7446 | -104.7919
26.7483 | -104.7937
26.752 | -104.793
26.7557 | -104.7939
26.7595 | -104.7958
26.7632 | -104.7965
26.7669 | -104.7972
26.7706 | -104.7978
26.7743 | -104.7992
26.778 | -104.7977
26.7817 | -104.7959
26.7854 | -104.7959
26.7891 | -104.7963
26.7929 | -104.7974
26.7966 | -104.7978
26.8003 | -104.799
26.804 | -104.8013
26.8077 | -104.8026
26.8114 | -104.8028
26.8151 | -104.8028
26.8188 | -104.803
26.8225 | -104.8061
26.8263 | -104.8106
26.83 | -104.8127
26.8337 | -104.8134
26.8374 | -104.8139
26.8411 | -104.8159
26.8448 | -104.8176
26.8485 | -104.8192
26.8522 | -104.8231
26.8559 | -104.8263
26.8597 | -104.8294
26.8634 | -104.8321
26.8671 | -104.8346
26.8708 | -104.8383
26.8745 | -104.8418
26.8782 | -104.8445
26.8819 | -104.8478
26.8856 | -104.8543
26.8894 | -104.8614
26.8931 | -104.8669
26.8968 | -104.8717
26.9005 | -104.8777
26.9042 | -104.8808
26.9079 | -104.8819
26.9116 | -104.8826
26.9153 | -104.8845
26.919 | -104.8848
26.9228 | -104.8869
26.9265 | -104.8887
26.9302 | -104.8893
26.9339 | -104.8877
26.9376 | -104.8875
26.9413 | -104.8881
26.945 | -104.8893
26.9487 | -104.8904
26.9524 | -104.8919
26.9562 | -104.8916
26.9599 | -104.8914
26.9636 | -104.8907
26.9673 | -104.8898
26.971 | -104.8872
26.9747 | -104.8844
26.9784 | -104.8811
26.9821 | -104.8777
26.9858 | -104.8746
26.9896 | -104.8705
26.9933 | -104.8678
26.997 | -104.8652
27.0007 | -104.8607
27.0044 | -104.8558
27.0081 | -104.8517
27.0118 | -104.8484
27.0155 | -104.8438
27.0193 | -104.8413
27.023 | -104.8377
27.0267 | -104.8317
27.0304 | -104.8257
27.0341 | -104.8221
27.0378 | -104.8183
27.0415 | -104.815
27.0452 | -104.81
27.0489 | -104.8052
27.0527 | -104.8002
27.0564 | -104.7953
27.0601 | -104.7912
27.0638 | -104.7879
27.0675 | -104.7834
27.0712 | -104.7787
27.0749 | -104.7748
27.0786 | -104.7706
27.0823 | -104.7663
27.0861 | -104.7619
27.0898 | -104.7563
27.0935 | -104.7521
27.0972 | -104.7485
27.1009 | -104.7445
27.1046 | -104.7402
27.1083 | -104.7355
27.112 | -104.7301
27.1157 | -104.7259
27.1195 | -104.723
27.1232 | -104.7189
27.1269 | -104.7128
27.1306 | -104.7093
27.1343 | -104.7036
27.138 | -104.6961
27.1417 | -104.6905
27.1454 | -104.6861
27.1492 | -104.6835
27.1529 | -104.6789
27.1566 | -104.6747
27.1603 | -104.6731
27.164 | -104.6704
27.1677 | -104.6663
27.1714 | -104.6617
27.1751 | -104.6565
27.1788 | -104.6532
27.1826 | -104.6498
27.1863 | -104.6457
27.19 | -104.6418
27.1937 | -104.6383
27.1974 | -104.636
27.2011 | -104.6335
27.2048 | -104.6279
27.2085 | -104.6214
27.2122 | -104.6174
27.216 | -104.6142
27.2197 | -104.6131
27.2234 | -104.6098
27.2271 | -104.6073
27.2308 | -104.6056
27.2345 | -104.6039
27.2382 | -104.6038
27.2419 | -104.6021
27.2457 | -104.601
27.2494 | -104.6014
27.2531 | -104.6011
27.2568 | -104.6
27.2605 | -104.5991
27.2642 | -104.5991
27.2679 | -104.5964
27.2716 | -104.5951
27.2753 | -104.5971
27.2791 | -104.5996
27.2828 | -104.6013
27.2865 | -104.6024
27.2902 | -104.6044
27.2939 | -104.6086
27.2976 | -104.6116
27.3013 | -104.6146
27.305 | -104.6145
27.3087 | -104.6158
27.3125 | -104.6165
27.3162 | -104.6179
27.3199 | -104.6207
27.3236 | -104.6235
27.3273 | -104.6255
27.331 | -104.6258
27.3347 | -104.6259
27.3384 | -104.6261
27.3421 | -104.6256
27.3459 | -104.6242
27.3496 | -104.625
27.3533 | -104.6261
27.357 | -104.6264
27.3607 | -104.6247
27.3644 | -104.6242
27.3681 | -104.6229
27.3718 | -104.6213
27.3756 | -104.6189
27.3793 | -104.6145
27.383 | -104.61
27.3867 | -104.6053
27.3904 | -104.6024
27.3941 | -104.5976
27.3978 | -104.5928
27.4015 | -104.5872
27.4052 | -104.581
27.409 | -104.575
27.4127 | -104.5684
27.4164 | -104.5621
27.4201 | -104.5541
27.4238 | -104.5466
27.4275 | -104.5379
27.4312 | -104.5296
27.4349 | -104.5217
27.4386 | -104.5138
27.4424 | -104.5052
27.4461 | -104.4979
27.4498 | -104.4902
27.4535 | -104.4829
27.4572 | -104.4751
27.4609 | -104.4685
27.4646 | -104.4622
27.4683 | -104.4555
27.472 | -104.449
27.4758 | -104.4427
27.4795 | -104.4336
27.4832 | -104.4261
27.4869 | -104.4204
27.4906 | -104.4139
27.4943 | -104.4088
27.498 | -104.4043
27.5017 | -104.3976
27.5055 | -104.3901
27.5092 | -104.3859
27.5129 | -104.3821
27.5166 | -104.3782
27.5203 | -104.3758
27.524 | -104.3744
27.5277 | -104.369
27.5314 | -104.3642
27.5351 | -104.3595
27.5389 | -104.3569
27.5426 | -104.3547
27.5463 | -104.351
27.55 | -104.349
27.5537 | -104.3482
27.5574 | -104.3489
27.5611 | -104.3489
27.5648 | -104.3471
27.5685 | -104.3479
27.5723 | -104.3498
27.576 | -104.3519
27.5797 | -104.3537
27.5834 | -104.3556
27.5871 | -104.3563
27.5908 | -104.3579
27.5945 | -104.3612
27.5982 | -104.3641
27.6019 | -104.3673
27.6057 | -104.3698
27.6094 | -104.3735
27.6131 | -104.3757
27.6168 | -104.3782
27.6205 | -104.3808
27.6242 | -104.383
27.6279 | -104.3855
27.6316 | -104.3887
27.6354 | -104.3925
27.6391 | -104.3959
27.6428 | -104.3994
27.6465 | -104.4038
27.6502 | -104.4075
27.6539 | -104.4138
27.6576 | -104.4185
27.6613 | -104.4208
27.665 | -104.4239
27.6688 | -104.4271
27.6725 | -104.4319
27.6762 | -104.4363
27.6799 | -104.4391
27.6836 | -104.4413
27.6873 | -104.4449
27.691 | -104.4508
27.6947 | -104.4565
27.6984 | -104.4617
27.7022 | -104.4661
27.7059 | -104.4675
27.7096 | -104.4705
27.7133 | -104.4732
27.717 | -104.476
27.7207 | -104.4784
27.7244 | -104.4801
27.7281 | -104.4829
27.7318 | -104.4868
27.7356 | -104.4911
27.7393 | -104.4923
27.743 | -104.4933
27.7467 | -104.4945
27.7504 | -104.4974
27.7541 | -104.4979
27.7578 | -104.4997
27.7615 | -104.4998
27.7653 | -104.5001
27.769 | -104.5022
27.7727 | -104.5044
27.7764 | -104.5057
27.7801 | -104.5052
27.7838 | -104.506
27.7875 | -104.5068
27.7912 | -104.5081
27.7949 | -104.5079
27.7987 | -104.5074
27.8024 | -104.5081
27.8061 | -104.5076
27.8098 | -104.5063
27.8135 | -104.5047
27.8172 | -104.5041
27.8209 | -104.5043
27.8246 | -104.5036
27.8283 | -104.5022
27.8321 | -104.5013
27.8358 | -104.5001
27.8395 | -104.4993
27.8432 | -104.4981
27.8469 | -104.4975
27.8506 | -104.498
27.8543 | -104.496
27.858 | -104.4926
27.8617 | -104.4905
27.8655 | -104.4886
27.8692 | -104.4885
27.8729 | -104.4897
27.8766 | -104.4879
27.8803 | -104.4863
27.884 | -104.4856
27.8877 | -104.4859
27.8914 | -104.4858
27.8952 | -104.4852
27.8989 | -104.4854
27.9026 | -104.485
27.9063 | -104.4863
27.91 | -104.4883
27.9137 | -104.489
27.9174 | -104.4895
27.9211 | -104.4916
27.9248 | -104.4895
27.9286 | -104.4806
27.9323 | -104.4726
27.936 | -104.4701
27.9397 | -104.4693
27.9434 | -104.4676
27.9471 | -104.4632
27.9508 | -104.4583
27.9545 | -104.4558
27.9582 | -104.4543
27.962 | -104.4493
27.9657 | -104.4437
27.9694 | -104.4397
27.9731 | -104.437
27.9768 | -104.4336
27.9805 | -104.4306
27.9842 | -104.4273
27.9879 | -104.4222
27.9916 | -104.415
27.9954 | -104.4064
27.9991 | -104.3952
28.0028 | -104.3859