## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 259 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 246
Hood | 02-92B/P | 208
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1003.2 | hPa |  | PRC PRO 01
Humidity | 57 | %/Hr
Internal temperature of calibration bench | 23.97 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1009 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.4 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.03 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.98 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3115 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 02/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 11/07/2018
Technical Manager | Anthony HUE | Approval |  | 16/07/2017
CEO | Robert GLEMOT | Approval |  | 16/07/2017
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4662 | 57.5253
0.012209 | -35.9974 | 38.4545
0.018313 | -34.9436 | 28.4974
0.024417 | -34.4925 | 22.1232
0.030521 | -34.388 | 18.34
0.036626 | -34.2466 | 15.2402
0.04273 | -34.1632 | 12.9225
0.048834 | -34.1111 | 11.3478
0.054938 | -34.0394 | 10.1461
0.061043 | -33.9635 | 9.0076
0.067147 | -33.9533 | 8.1353
0.073251 | -33.9548 | 7.38
0.079355 | -33.9737 | 6.8195
0.08546 | -33.9921 | 6.3582
0.091564 | -33.9883 | 5.9826
0.097668 | -33.9539 | 5.7301
0.10377 | -33.9064 | 5.3861
0.10988 | -33.8755 | 5.0245
0.11598 | -33.8662 | 4.6635
0.12209 | -33.8828 | 4.381
0.12819 | -33.9152 | 4.1437
0.13429 | -33.9468 | 3.8999
0.1404 | -33.9637 | 3.6262
0.1465 | -33.9792 | 3.3677
0.15261 | -33.9861 | 3.1289
0.15871 | -33.9826 | 2.9339
0.16482 | -33.9683 | 2.7773
0.17092 | -33.9444 | 2.7165
0.17702 | -33.9259 | 2.62
0.18313 | -33.9091 | 2.5376
0.18923 | -33.897 | 2.4314
0.19534 | -33.8903 | 2.3272
0.20144 | -33.8893 | 2.2129
0.20754 | -33.8919 | 2.1377
0.21365 | -33.8943 | 2.0617
0.21975 | -33.8986 | 1.9961
0.22586 | -33.9065 | 1.9302
0.23196 | -33.915 | 1.8575
0.23807 | -33.9236 | 1.7803
0.24417 | -33.9258 | 1.6882
0.25027 | -33.9293 | 1.6047
0.25638 | -33.9293 | 1.5137
0.26248 | -33.9275 | 1.4275
0.26859 | -33.9251 | 1.345
0.27469 | -33.9244 | 1.2589
0.2808 | -33.9301 | 1.1941
0.2869 | -33.9345 | 1.1254
0.293 | -33.9398 | 1.0677
0.29911 | -33.946 | 1.0284
0.30521 | -33.9519 | 0.99089
0.31132 | -33.9558 | 0.97643
0.31742 | -33.9601 | 0.9537
0.32353 | -33.9634 | 0.93523
0.32963 | -33.9663 | 0.92002
0.33573 | -33.9681 | 0.90714
0.34184 | -33.9706 | 0.891
0.34794 | -33.9737 | 0.85637
0.35405 | -33.977 | 0.82675
0.36015 | -33.9805 | 0.79358
0.36626 | -33.9828 | 0.75685
0.37236 | -33.9852 | 0.70723
0.37846 | -33.9873 | 0.65001
0.38457 | -33.9863 | 0.59054
0.39067 | -33.9869 | 0.54015
0.39678 | -33.9852 | 0.48354
0.40288 | -33.9829 | 0.42997
0.40899 | -33.9791 | 0.37771
0.41509 | -33.9747 | 0.32526
0.42119 | -33.9697 | 0.27782
0.4273 | -33.9661 | 0.2287
0.4334 | -33.9627 | 0.17873
0.43951 | -33.9597 | 0.12968
0.44561 | -33.9582 | 0.083841
0.45172 | -33.9598 | 0.045964
0.45782 | -33.9596 | 0.0037651
0.46392 | -33.9609 | -0.035052
0.47003 | -33.9626 | -0.070377
0.47613 | -33.9646 | -0.10272
0.48224 | -33.9668 | -0.13359
0.48834 | -33.9692 | -0.15556
0.49445 | -33.9713 | -0.1808
0.50055 | -33.9738 | -0.20363
0.50665 | -33.9755 | -0.22236
0.51276 | -33.9767 | -0.24257
0.51886 | -33.9778 | -0.26026
0.52497 | -33.9769 | -0.26691
0.53107 | -33.9773 | -0.27744
0.53717 | -33.9777 | -0.28885
0.54328 | -33.9775 | -0.29411
0.54938 | -33.9774 | -0.30228
0.55549 | -33.977 | -0.31128
0.56159 | -33.9769 | -0.32207
0.5677 | -33.9765 | -0.33233
0.5738 | -33.9761 | -0.34543
0.5799 | -33.9762 | -0.35737
0.58601 | -33.9767 | -0.37357
0.59211 | -33.9778 | -0.3958
0.59822 | -33.9785 | -0.41228
0.60432 | -33.9795 | -0.43367
0.61043 | -33.9806 | -0.45464
0.61653 | -33.9818 | -0.47602
0.62263 | -33.9829 | -0.49895
0.62874 | -33.9843 | -0.5192
0.63484 | -33.9855 | -0.54018
0.64095 | -33.9867 | -0.5611
0.64705 | -33.9879 | -0.58095
0.65316 | -33.9889 | -0.60046
0.65926 | -33.9899 | -0.62299
0.66536 | -33.9894 | -0.6414
0.67147 | -33.9896 | -0.65772
0.67757 | -33.9902 | -0.67566
0.68368 | -33.99 | -0.69573
0.68978 | -33.99 | -0.71378
0.69589 | -33.9896 | -0.73215
0.70199 | -33.9888 | -0.75405
0.70809 | -33.9882 | -0.77424
0.7142 | -33.9877 | -0.79354
0.7203 | -33.9869 | -0.81309
0.72641 | -33.986 | -0.83399
0.73251 | -33.9858 | -0.85476
0.73862 | -33.9858 | -0.87908
0.74472 | -33.986 | -0.89843
0.75082 | -33.9856 | -0.91954
0.75693 | -33.9857 | -0.94306
0.76303 | -33.9856 | -0.96558
0.76914 | -33.9862 | -0.99263
0.77524 | -33.9864 | -1.0175
0.78135 | -33.9868 | -1.0409
0.78745 | -33.9875 | -1.0658
0.79355 | -33.988 | -1.0904
0.79966 | -33.989 | -1.1162
0.80576 | -33.9902 | -1.1466
0.81187 | -33.9906 | -1.1761
0.81797 | -33.9912 | -1.2068
0.82408 | -33.9919 | -1.234
0.83018 | -33.9925 | -1.2651
0.83628 | -33.9934 | -1.2955
0.84239 | -33.9933 | -1.3235
0.84849 | -33.9933 | -1.3516
0.8546 | -33.9938 | -1.3792
0.8607 | -33.9939 | -1.4084
0.86681 | -33.994 | -1.4367
0.87291 | -33.9936 | -1.4623
0.87901 | -33.9931 | -1.4845
0.88512 | -33.993 | -1.5075
0.89122 | -33.9929 | -1.5307
0.89733 | -33.9926 | -1.5499
0.90343 | -33.9923 | -1.571
0.90953 | -33.9925 | -1.5858
0.91564 | -33.992 | -1.6017
0.92174 | -33.9916 | -1.6158
0.92785 | -33.9912 | -1.6292
0.93395 | -33.9912 | -1.6408
0.94006 | -33.9912 | -1.6543
0.94616 | -33.991 | -1.6588
0.95226 | -33.9911 | -1.6702
0.95837 | -33.9912 | -1.6818
0.96447 | -33.9916 | -1.6919
0.97058 | -33.9916 | -1.7017
0.97668 | -33.9914 | -1.7096
0.98279 | -33.9914 | -1.7132
0.98889 | -33.9917 | -1.7179
0.99499 | -33.9917 | -1.7264
1.0011 | -33.992 | -1.7312
1.0072 | -33.9919 | -1.7349
1.0133 | -33.9918 | -1.7387
1.0194 | -33.9918 | -1.7442
1.0255 | -33.9915 | -1.7511
1.0316 | -33.9909 | -1.7552
1.0377 | -33.9903 | -1.7611
1.0438 | -33.9896 | -1.7656
1.0499 | -33.988 | -1.7721
1.056 | -33.9869 | -1.7771
1.0621 | -33.9857 | -1.7829
1.0682 | -33.9852 | -1.7879
1.0743 | -33.9841 | -1.7958
1.0805 | -33.9834 | -1.8027
1.0866 | -33.9818 | -1.8136
1.0927 | -33.9804 | -1.8211
1.0988 | -33.9796 | -1.8317
1.1049 | -33.9781 | -1.8394
1.111 | -33.9765 | -1.8506
1.1171 | -33.9751 | -1.86
1.1232 | -33.9737 | -1.8742
1.1293 | -33.9726 | -1.8862
1.1354 | -33.9716 | -1.8986
1.1415 | -33.9704 | -1.9126
1.1476 | -33.969 | -1.9267
1.1537 | -33.9676 | -1.9399
1.1598 | -33.9675 | -1.9581
1.1659 | -33.9669 | -1.9701
1.172 | -33.9662 | -1.9834
1.1781 | -33.9655 | -1.9974
1.1842 | -33.9649 | -2.0129
1.1903 | -33.9656 | -2.0279
1.1964 | -33.9655 | -2.0423
1.2025 | -33.9654 | -2.0575
1.2086 | -33.9659 | -2.0731
1.2147 | -33.9661 | -2.088
1.2209 | -33.9668 | -2.1013
1.227 | -33.9687 | -2.1162
1.2331 | -33.9697 | -2.1304
1.2392 | -33.9706 | -2.1461
1.2453 | -33.9718 | -2.158
1.2514 | -33.9731 | -2.1699
1.2575 | -33.9747 | -2.1838
1.2636 | -33.977 | -2.195
1.2697 | -33.9788 | -2.2066
1.2758 | -33.9803 | -2.2205
1.2819 | -33.9821 | -2.2339
1.288 | -33.9838 | -2.2475
1.2941 | -33.9859 | -2.2579
1.3002 | -33.9879 | -2.266
1.3063 | -33.9895 | -2.2758
1.3124 | -33.9914 | -2.2857
1.3185 | -33.9933 | -2.2961
1.3246 | -33.9952 | -2.3067
1.3307 | -33.9973 | -2.316
1.3368 | -33.9992 | -2.3234
1.3429 | -34.0008 | -2.332
1.349 | -34.0023 | -2.3404
1.3551 | -34.004 | -2.3511
1.3613 | -34.0056 | -2.3588
1.3674 | -34.0068 | -2.3671
1.3735 | -34.0081 | -2.3751
1.3796 | -34.0094 | -2.3823
1.3857 | -34.0104 | -2.3892
1.3918 | -34.0114 | -2.3988
1.3979 | -34.0128 | -2.4071
1.404 | -34.0134 | -2.417
1.4101 | -34.0143 | -2.4259
1.4162 | -34.0152 | -2.4343
1.4223 | -34.016 | -2.4445
1.4284 | -34.0168 | -2.4528
1.4345 | -34.0174 | -2.4629
1.4406 | -34.0178 | -2.4741
1.4467 | -34.0185 | -2.4835
1.4528 | -34.0192 | -2.4927
1.4589 | -34.02 | -2.5031
1.465 | -34.0202 | -2.5108
1.4711 | -34.0212 | -2.5198
1.4772 | -34.0217 | -2.5319
1.4833 | -34.0223 | -2.5441
1.4894 | -34.023 | -2.5555
1.4955 | -34.0237 | -2.567
1.5016 | -34.0244 | -2.579
1.5078 | -34.0249 | -2.5936
1.5139 | -34.0255 | -2.6063
1.52 | -34.026 | -2.6194
1.5261 | -34.0264 | -2.6311
1.5322 | -34.0266 | -2.6438
1.5383 | -34.0269 | -2.6559
1.5444 | -34.0276 | -2.6693
1.5505 | -34.028 | -2.6826
1.5566 | -34.0285 | -2.6954
1.5627 | -34.0291 | -2.7088
1.5688 | -34.0299 | -2.722
1.5749 | -34.0301 | -2.7343
1.581 | -34.0305 | -2.751
1.5871 | -34.031 | -2.7666
1.5932 | -34.0314 | -2.7795
1.5993 | -34.0318 | -2.7963
1.6054 | -34.0321 | -2.8113
1.6115 | -34.0325 | -2.8282
1.6176 | -34.0328 | -2.8446
1.6237 | -34.0331 | -2.8591
1.6298 | -34.0333 | -2.8744
1.6359 | -34.0337 | -2.889
1.642 | -34.0342 | -2.9034
1.6482 | -34.0345 | -2.9204
1.6543 | -34.035 | -2.9372
1.6604 | -34.0354 | -2.9519
1.6665 | -34.0355 | -2.9668
1.6726 | -34.0358 | -2.983
1.6787 | -34.0363 | -3
1.6848 | -34.0361 | -3.0164
1.6909 | -34.0365 | -3.0327
1.697 | -34.0367 | -3.0488
1.7031 | -34.0367 | -3.0645
1.7092 | -34.0367 | -3.0795
1.7153 | -34.0369 | -3.0956
1.7214 | -34.0365 | -3.111
1.7275 | -34.0363 | -3.1268
1.7336 | -34.0361 | -3.1431
1.7397 | -34.0361 | -3.1584
1.7458 | -34.0361 | -3.175
1.7519 | -34.0359 | -3.1911
1.758 | -34.0358 | -3.2061
1.7641 | -34.0359 | -3.2205
1.7702 | -34.0356 | -3.2363
1.7763 | -34.0357 | -3.25
1.7824 | -34.0355 | -3.2662
1.7885 | -34.0353 | -3.2804
1.7947 | -34.0353 | -3.2953
1.8008 | -34.0352 | -3.3082
1.8069 | -34.0351 | -3.3219
1.813 | -34.0351 | -3.3361
1.8191 | -34.0348 | -3.3504
1.8252 | -34.0346 | -3.3644
1.8313 | -34.0346 | -3.3788
1.8374 | -34.0346 | -3.3924
1.8435 | -34.0347 | -3.4046
1.8496 | -34.0344 | -3.4192
1.8557 | -34.0343 | -3.434
1.8618 | -34.0342 | -3.4458
1.8679 | -34.0342 | -3.4589
1.874 | -34.0342 | -3.4721
1.8801 | -34.0341 | -3.4851
1.8862 | -34.0342 | -3.498
1.8923 | -34.0342 | -3.5102
1.8984 | -34.0338 | -3.5209
1.9045 | -34.0337 | -3.5321
1.9106 | -34.0335 | -3.5439
1.9167 | -34.0332 | -3.5553
1.9228 | -34.0331 | -3.5662
1.9289 | -34.0328 | -3.5751
1.9351 | -34.0329 | -3.5842
1.9412 | -34.0329 | -3.5946
1.9473 | -34.033 | -3.6066
1.9534 | -34.0329 | -3.6178
1.9595 | -34.0327 | -3.6293
1.9656 | -34.0322 | -3.6401
1.9717 | -34.032 | -3.6506
1.9778 | -34.032 | -3.6621
1.9839 | -34.032 | -3.6736
1.99 | -34.032 | -3.6834
1.9961 | -34.0318 | -3.6939
2.0022 | -34.0317 | -3.7039
2.0083 | -34.0314 | -3.7139
2.0144 | -34.0314 | -3.7237
2.0205 | -34.0314 | -3.7356
2.0266 | -34.0315 | -3.7471
2.0327 | -34.0315 | -3.7583
2.0388 | -34.0318 | -3.7721
2.0449 | -34.0318 | -3.7824
2.051 | -34.0319 | -3.7924
2.0571 | -34.0317 | -3.8037
2.0632 | -34.0316 | -3.8143
2.0693 | -34.0316 | -3.8255
2.0754 | -34.0318 | -3.8372
2.0816 | -34.032 | -3.8492
2.0877 | -34.0319 | -3.8603
2.0938 | -34.0321 | -3.8727
2.0999 | -34.0321 | -3.8842
2.106 | -34.0323 | -3.8961
2.1121 | -34.0323 | -3.9084
2.1182 | -34.0324 | -3.9184
2.1243 | -34.0326 | -3.9289
2.1304 | -34.0327 | -3.9392
2.1365 | -34.033 | -3.9507
2.1426 | -34.0331 | -3.964
2.1487 | -34.0336 | -3.9757
2.1548 | -34.034 | -3.9887
2.1609 | -34.0346 | -4.0007
2.167 | -34.035 | -4.0132
2.1731 | -34.0353 | -4.0251
2.1792 | -34.036 | -4.0364
2.1853 | -34.0365 | -4.0492
2.1914 | -34.037 | -4.0614
2.1975 | -34.0376 | -4.0735
2.2036 | -34.038 | -4.0867
2.2097 | -34.0384 | -4.0987
2.2158 | -34.0393 | -4.109
2.222 | -34.0399 | -4.1191
2.2281 | -34.0402 | -4.1299
2.2342 | -34.0407 | -4.143
2.2403 | -34.0411 | -4.1531
2.2464 | -34.0419 | -4.1641
2.2525 | -34.0423 | -4.1749
2.2586 | -34.0432 | -4.1848
2.2647 | -34.044 | -4.1941
2.2708 | -34.0446 | -4.2058
2.2769 | -34.0455 | -4.2173
2.283 | -34.0459 | -4.2268
2.2891 | -34.0465 | -4.2371
2.2952 | -34.047 | -4.2469
2.3013 | -34.0477 | -4.2571
2.3074 | -34.0484 | -4.2668
2.3135 | -34.0488 | -4.2781
2.3196 | -34.0494 | -4.2867
2.3257 | -34.0498 | -4.2973
2.3318 | -34.0501 | -4.3064
2.3379 | -34.0506 | -4.3178
2.344 | -34.0509 | -4.3287
2.3501 | -34.0514 | -4.3395
2.3562 | -34.0517 | -4.3479
2.3623 | -34.052 | -4.3566
2.3685 | -34.0522 | -4.3657
2.3746 | -34.0525 | -4.3762
2.3807 | -34.0529 | -4.3852
2.3868 | -34.0531 | -4.3943
2.3929 | -34.0535 | -4.4037
2.399 | -34.0538 | -4.4118
2.4051 | -34.0543 | -4.4213
2.4112 | -34.0546 | -4.4304
2.4173 | -34.0549 | -4.4393
2.4234 | -34.0549 | -4.4479
2.4295 | -34.0552 | -4.4566
2.4356 | -34.0553 | -4.466
2.4417 | -34.0553 | -4.4761
2.4478 | -34.0554 | -4.4857
2.4539 | -34.0557 | -4.4937
2.46 | -34.0554 | -4.5035
2.4661 | -34.0554 | -4.5129
2.4722 | -34.0557 | -4.5216
2.4783 | -34.0558 | -4.5301
2.4844 | -34.0558 | -4.5387
2.4905 | -34.0557 | -4.548
2.4966 | -34.0554 | -4.557
2.5027 | -34.0554 | -4.5651
2.5089 | -34.0553 | -4.5737
2.515 | -34.0551 | -4.5831
2.5211 | -34.055 | -4.5915
2.5272 | -34.0547 | -4.6006
2.5333 | -34.0545 | -4.6102
2.5394 | -34.0545 | -4.6193
2.5455 | -34.0543 | -4.6288
2.5516 | -34.0543 | -4.6376
2.5577 | -34.0542 | -4.6477
2.5638 | -34.054 | -4.6581
2.5699 | -34.0538 | -4.6685
2.576 | -34.0535 | -4.679
2.5821 | -34.0534 | -4.689
2.5882 | -34.0532 | -4.6987
2.5943 | -34.053 | -4.7086
2.6004 | -34.0528 | -4.7186
2.6065 | -34.0527 | -4.7272
2.6126 | -34.0526 | -4.7364
2.6187 | -34.0525 | -4.7452
2.6248 | -34.0523 | -4.7543
2.6309 | -34.0522 | -4.7657
2.637 | -34.0522 | -4.777
2.6431 | -34.0522 | -4.7866
2.6492 | -34.052 | -4.7967
2.6554 | -34.052 | -4.8071
2.6615 | -34.0518 | -4.817
2.6676 | -34.0518 | -4.8276
2.6737 | -34.0519 | -4.8385
2.6798 | -34.0519 | -4.8485
2.6859 | -34.0519 | -4.8591
2.692 | -34.0518 | -4.869
2.6981 | -34.0518 | -4.8802
2.7042 | -34.0519 | -4.8923
2.7103 | -34.0518 | -4.9025
2.7164 | -34.0521 | -4.9133
2.7225 | -34.0522 | -4.9247
2.7286 | -34.0522 | -4.936
2.7347 | -34.0523 | -4.9461
2.7408 | -34.0526 | -4.9554
2.7469 | -34.0525 | -4.9669
2.753 | -34.0528 | -4.9776
2.7591 | -34.053 | -4.9897
2.7652 | -34.053 | -5.0013
2.7713 | -34.053 | -5.0113
2.7774 | -34.0534 | -5.0228
2.7835 | -34.0535 | -5.0335
2.7896 | -34.0536 | -5.0429
2.7958 | -34.054 | -5.0534
2.8019 | -34.0545 | -5.0632
2.808 | -34.0548 | -5.0752
2.8141 | -34.0552 | -5.0882
2.8202 | -34.0556 | -5.1005
2.8263 | -34.056 | -5.1115
2.8324 | -34.0563 | -5.124
2.8385 | -34.0567 | -5.1362
2.8446 | -34.0573 | -5.148
2.8507 | -34.0576 | -5.1601
2.8568 | -34.0581 | -5.1714
2.8629 | -34.0585 | -5.1828
2.869 | -34.059 | -5.1942
2.8751 | -34.0594 | -5.2062
2.8812 | -34.06 | -5.2196
2.8873 | -34.0603 | -5.231
2.8934 | -34.0607 | -5.243
2.8995 | -34.0611 | -5.2551
2.9056 | -34.0614 | -5.2659
2.9117 | -34.0618 | -5.2776
2.9178 | -34.0625 | -5.2903
2.9239 | -34.0629 | -5.3037
2.93 | -34.0634 | -5.3173
2.9361 | -34.064 | -5.3288
2.9423 | -34.0645 | -5.3408
2.9484 | -34.065 | -5.3544
2.9545 | -34.0657 | -5.3677
2.9606 | -34.0662 | -5.3811
2.9667 | -34.0668 | -5.3931
2.9728 | -34.0671 | -5.4054
2.9789 | -34.0676 | -5.4182
2.985 | -34.068 | -5.432
2.9911 | -34.0685 | -5.4447
2.9972 | -34.0691 | -5.4574
3.0033 | -34.0697 | -5.4706
3.0094 | -34.0702 | -5.4837
3.0155 | -34.0708 | -5.4963
3.0216 | -34.0713 | -5.5098
3.0277 | -34.0717 | -5.5236
3.0338 | -34.0723 | -5.5372
3.0399 | -34.0729 | -5.551
3.046 | -34.0732 | -5.5641
3.0521 | -34.0738 | -5.5766
3.0582 | -34.0741 | -5.5906
3.0643 | -34.0746 | -5.6037
3.0704 | -34.0749 | -5.6171
3.0765 | -34.0753 | -5.6294
3.0827 | -34.0756 | -5.6426
3.0888 | -34.0759 | -5.6555
3.0949 | -34.0763 | -5.6685
3.101 | -34.0767 | -5.6827
3.1071 | -34.077 | -5.6969
3.1132 | -34.0775 | -5.711
3.1193 | -34.0779 | -5.724
3.1254 | -34.0783 | -5.7363
3.1315 | -34.0787 | -5.7507
3.1376 | -34.079 | -5.7636
3.1437 | -34.0793 | -5.7772
3.1498 | -34.0797 | -5.7914
3.1559 | -34.08 | -5.8047
3.162 | -34.08 | -5.8189
3.1681 | -34.0801 | -5.8331
3.1742 | -34.0805 | -5.8469
3.1803 | -34.0808 | -5.8601
3.1864 | -34.0812 | -5.8725
3.1925 | -34.0815 | -5.8862
3.1986 | -34.0816 | -5.899
3.2047 | -34.0817 | -5.9131
3.2108 | -34.082 | -5.9249
3.2169 | -34.0824 | -5.9387
3.223 | -34.0823 | -5.9521
3.2292 | -34.0823 | -5.9661
3.2353 | -34.0822 | -5.9757
3.2414 | -34.0824 | -5.9888
3.2475 | -34.0823 | -5.9998
3.2536 | -34.0826 | -6.0136
3.2597 | -34.0828 | -6.0249
3.2658 | -34.083 | -6.0357
3.2719 | -34.0831 | -6.0468
3.278 | -34.0832 | -6.0591
3.2841 | -34.0833 | -6.071
3.2902 | -34.0834 | -6.0836
3.2963 | -34.0836 | -6.0954
3.3024 | -34.0834 | -6.1056
3.3085 | -34.0837 | -6.1182
3.3146 | -34.0836 | -6.1292
3.3207 | -34.0838 | -6.1414
3.3268 | -34.0839 | -6.1537
3.3329 | -34.084 | -6.1635
3.339 | -34.0841 | -6.1733
3.3451 | -34.0841 | -6.1848
3.3512 | -34.0842 | -6.1943
3.3573 | -34.0844 | -6.2057
3.3634 | -34.0845 | -6.2172
3.3696 | -34.0844 | -6.2292
3.3757 | -34.0844 | -6.2391
3.3818 | -34.0847 | -6.2509
3.3879 | -34.0846 | -6.2613
3.394 | -34.0846 | -6.2719
3.4001 | -34.0846 | -6.2828
3.4062 | -34.0846 | -6.2931
3.4123 | -34.0847 | -6.3029
3.4184 | -34.0847 | -6.314
3.4245 | -34.0848 | -6.3246
3.4306 | -34.0847 | -6.3344
3.4367 | -34.0849 | -6.3442
3.4428 | -34.0851 | -6.3535
3.4489 | -34.0851 | -6.3646
3.455 | -34.085 | -6.3751
3.4611 | -34.085 | -6.3863
3.4672 | -34.085 | -6.3977
3.4733 | -34.085 | -6.4079
3.4794 | -34.085 | -6.4185
3.4855 | -34.0853 | -6.4281
3.4916 | -34.0854 | -6.4379
3.4977 | -34.0853 | -6.4489
3.5038 | -34.0855 | -6.459
3.5099 | -34.0855 | -6.4686
3.5161 | -34.0856 | -6.4782
3.5222 | -34.0857 | -6.4883
3.5283 | -34.0857 | -6.4984
3.5344 | -34.0857 | -6.5092
3.5405 | -34.0858 | -6.5192
3.5466 | -34.0859 | -6.5291
3.5527 | -34.086 | -6.5398
3.5588 | -34.0861 | -6.5496
3.5649 | -34.0861 | -6.5599
3.571 | -34.0862 | -6.5701
3.5771 | -34.0863 | -6.5809
3.5832 | -34.0864 | -6.5906
3.5893 | -34.0866 | -6.6005
3.5954 | -34.0867 | -6.6111
3.6015 | -34.0868 | -6.6215
3.6076 | -34.0871 | -6.6312
3.6137 | -34.0871 | -6.6411
3.6198 | -34.0875 | -6.6502
3.6259 | -34.0875 | -6.661
3.632 | -34.0877 | -6.6717
3.6381 | -34.0879 | -6.6816
3.6442 | -34.0881 | -6.6921
3.6503 | -34.0883 | -6.7013
3.6565 | -34.0886 | -6.713
3.6626 | -34.0887 | -6.7238
3.6687 | -34.089 | -6.7344
3.6748 | -34.0893 | -6.7444
3.6809 | -34.0896 | -6.7549
3.687 | -34.0898 | -6.7662
3.6931 | -34.0902 | -6.7762
3.6992 | -34.0904 | -6.7869
3.7053 | -34.0906 | -6.7971
3.7114 | -34.0908 | -6.8084
3.7175 | -34.0911 | -6.8183
3.7236 | -34.0913 | -6.8272
3.7297 | -34.0914 | -6.838
3.7358 | -34.0917 | -6.8497
3.7419 | -34.0918 | -6.8613
3.748 | -34.092 | -6.8721
3.7541 | -34.0921 | -6.8831
3.7602 | -34.0925 | -6.894
3.7663 | -34.0927 | -6.9042
3.7724 | -34.093 | -6.9153
3.7785 | -34.0933 | -6.9257
3.7846 | -34.0935 | -6.9363
3.7907 | -34.0939 | -6.9474
3.7969 | -34.0941 | -6.9589
3.803 | -34.0945 | -6.969
3.8091 | -34.0948 | -6.9794
3.8152 | -34.095 | -6.9906
3.8213 | -34.0952 | -7.0023
3.8274 | -34.0956 | -7.0128
3.8335 | -34.096 | -7.0244
3.8396 | -34.0964 | -7.0359
3.8457 | -34.0968 | -7.0463
3.8518 | -34.0973 | -7.0572
3.8579 | -34.0976 | -7.0678
3.864 | -34.098 | -7.0785
3.8701 | -34.0984 | -7.0896
3.8762 | -34.0987 | -7.1008
3.8823 | -34.0989 | -7.1115
3.8884 | -34.0993 | -7.1219
3.8945 | -34.0998 | -7.1332
3.9006 | -34.1002 | -7.1448
3.9067 | -34.1005 | -7.1567
3.9128 | -34.1009 | -7.168
3.9189 | -34.1013 | -7.1787
3.925 | -34.1015 | -7.1896
3.9311 | -34.1019 | -7.1996
3.9372 | -34.1023 | -7.2112
3.9434 | -34.1026 | -7.2217
3.9495 | -34.103 | -7.2334
3.9556 | -34.1035 | -7.2446
3.9617 | -34.1038 | -7.2555
3.9678 | -34.1042 | -7.2663
3.9739 | -34.1046 | -7.2775
3.98 | -34.1049 | -7.2894
3.9861 | -34.1052 | -7.2998
3.9922 | -34.1054 | -7.3115
3.9983 | -34.1058 | -7.3216
4.0044 | -34.1061 | -7.3314
4.0105 | -34.1065 | -7.3423
4.0166 | -34.107 | -7.3534
4.0227 | -34.1074 | -7.3638
4.0288 | -34.1076 | -7.3752
4.0349 | -34.1079 | -7.3852
4.041 | -34.1082 | -7.3961
4.0471 | -34.1085 | -7.4064
4.0532 | -34.1089 | -7.4176
4.0593 | -34.1092 | -7.4292
4.0654 | -34.1095 | -7.44
4.0715 | -34.1099 | -7.4515
4.0776 | -34.1102 | -7.4619
4.0838 | -34.1105 | -7.4733
4.0899 | -34.1109 | -7.4846
4.096 | -34.1112 | -7.4951
4.1021 | -34.1115 | -7.5059
4.1082 | -34.1119 | -7.5162
4.1143 | -34.1122 | -7.5262
4.1204 | -34.1126 | -7.537
4.1265 | -34.113 | -7.5474
4.1326 | -34.1134 | -7.5577
4.1387 | -34.1138 | -7.569
4.1448 | -34.1141 | -7.58
4.1509 | -34.1143 | -7.5908
4.157 | -34.1146 | -7.602
4.1631 | -34.1149 | -7.6133
4.1692 | -34.1152 | -7.624
4.1753 | -34.1156 | -7.6349
4.1814 | -34.1158 | -7.6448
4.1875 | -34.1162 | -7.6562
4.1936 | -34.1164 | -7.6676
4.1997 | -34.1167 | -7.6777
4.2058 | -34.117 | -7.6893
4.2119 | -34.1173 | -7.6991
4.218 | -34.1175 | -7.7092
4.2241 | -34.1179 | -7.7189
4.2303 | -34.1183 | -7.7302
4.2364 | -34.1186 | -7.7416
4.2425 | -34.1189 | -7.7521
4.2486 | -34.1191 | -7.7633
4.2547 | -34.1192 | -7.7726
4.2608 | -34.1195 | -7.7826
4.2669 | -34.1198 | -7.793
4.273 | -34.1201 | -7.8038
4.2791 | -34.1203 | -7.8146
4.2852 | -34.1205 | -7.8249
4.2913 | -34.1206 | -7.8342
4.2974 | -34.1208 | -7.8452
4.3035 | -34.1211 | -7.8561
4.3096 | -34.1212 | -7.8666
4.3157 | -34.1215 | -7.8766
4.3218 | -34.1217 | -7.8866
4.3279 | -34.1219 | -7.897
4.334 | -34.1221 | -7.9069
4.3401 | -34.1223 | -7.9168
4.3462 | -34.1226 | -7.9277
4.3523 | -34.1229 | -7.9374
4.3584 | -34.1231 | -7.946
4.3645 | -34.1232 | -7.9557
4.3707 | -34.1235 | -7.965
4.3768 | -34.1237 | -7.9754
4.3829 | -34.1238 | -7.9858
4.389 | -34.1241 | -7.9956
4.3951 | -34.1243 | -8.0052
4.4012 | -34.1245 | -8.0151
4.4073 | -34.1247 | -8.0252
4.4134 | -34.1249 | -8.0354
4.4195 | -34.1252 | -8.0451
4.4256 | -34.1255 | -8.0548
4.4317 | -34.1257 | -8.0643
4.4378 | -34.1258 | -8.0739
4.4439 | -34.126 | -8.0834
4.45 | -34.1261 | -8.0934
4.4561 | -34.1262 | -8.1033
4.4622 | -34.1264 | -8.113
4.4683 | -34.1265 | -8.1223
4.4744 | -34.1267 | -8.1325
4.4805 | -34.127 | -8.1417
4.4866 | -34.1272 | -8.151
4.4927 | -34.1274 | -8.1613
4.4988 | -34.1275 | -8.1707
4.5049 | -34.1278 | -8.1799
4.511 | -34.128 | -8.189
4.5172 | -34.1283 | -8.1989
4.5233 | -34.1285 | -8.2095
4.5294 | -34.1287 | -8.2189
4.5355 | -34.1289 | -8.2284
4.5416 | -34.1292 | -8.2376
4.5477 | -34.1293 | -8.2469
4.5538 | -34.1295 | -8.2557
4.5599 | -34.1299 | -8.2653
4.566 | -34.13 | -8.2751
4.5721 | -34.1303 | -8.2844
4.5782 | -34.1305 | -8.2943
4.5843 | -34.1307 | -8.3039
4.5904 | -34.1311 | -8.3133
4.5965 | -34.1313 | -8.3221
4.6026 | -34.1316 | -8.3328
4.6087 | -34.1318 | -8.3419
4.6148 | -34.132 | -8.3526
4.6209 | -34.1324 | -8.3611
4.627 | -34.1327 | -8.3712
4.6331 | -34.1328 | -8.3802
4.6392 | -34.133 | -8.3895
4.6453 | -34.1332 | -8.399
4.6514 | -34.1335 | -8.4086
4.6576 | -34.1335 | -8.4191
4.6637 | -34.1338 | -8.4269
4.6698 | -34.134 | -8.4355
4.6759 | -34.1341 | -8.4454
4.682 | -34.1343 | -8.4546
4.6881 | -34.1344 | -8.4635
4.6942 | -34.1347 | -8.4728
4.7003 | -34.1349 | -8.4824
4.7064 | -34.1352 | -8.4917
4.7125 | -34.1355 | -8.5013
4.7186 | -34.1356 | -8.5098
4.7247 | -34.136 | -8.5195
4.7308 | -34.1363 | -8.5289
4.7369 | -34.1364 | -8.5385
4.743 | -34.1366 | -8.549
4.7491 | -34.1368 | -8.5587
4.7552 | -34.1369 | -8.568
4.7613 | -34.1371 | -8.5777
4.7674 | -34.1374 | -8.5864
4.7735 | -34.1375 | -8.5966
4.7796 | -34.1377 | -8.6065
4.7857 | -34.1378 | -8.6162
4.7918 | -34.1381 | -8.6249
4.7979 | -34.1384 | -8.6341
4.8041 | -34.1385 | -8.6435
4.8102 | -34.1389 | -8.6526
4.8163 | -34.1391 | -8.6622
4.8224 | -34.1392 | -8.6709
4.8285 | -34.1395 | -8.6801
4.8346 | -34.1398 | -8.6897
4.8407 | -34.14 | -8.6991
4.8468 | -34.1402 | -8.7089
4.8529 | -34.1404 | -8.7188
4.859 | -34.1406 | -8.7284
4.8651 | -34.1408 | -8.7379
4.8712 | -34.1411 | -8.7473
4.8773 | -34.1413 | -8.757
4.8834 | -34.1415 | -8.7668
4.8895 | -34.1418 | -8.7769
4.8956 | -34.1419 | -8.787
4.9017 | -34.1421 | -8.7965
4.9078 | -34.1423 | -8.8065
4.9139 | -34.1426 | -8.8152
4.92 | -34.1428 | -8.8243
4.9261 | -34.1431 | -8.8332
4.9322 | -34.1435 | -8.8425
4.9383 | -34.1438 | -8.8516
4.9445 | -34.1441 | -8.8618
4.9506 | -34.1443 | -8.8707
4.9567 | -34.1445 | -8.8801
4.9628 | -34.1448 | -8.8896
4.9689 | -34.145 | -8.8995
4.975 | -34.1452 | -8.909
4.9811 | -34.1455 | -8.9188
4.9872 | -34.1457 | -8.9287
4.9933 | -34.1461 | -8.9385
4.9994 | -34.1464 | -8.9485
5.0055 | -34.1467 | -8.9582
5.0116 | -34.1471 | -8.9685
5.0177 | -34.1474 | -8.9787
5.0238 | -34.1477 | -8.9893
5.0299 | -34.1482 | -9.0004
5.036 | -34.1485 | -9.01
5.0421 | -34.1486 | -9.0198
5.0482 | -34.1489 | -9.0301
5.0543 | -34.1492 | -9.0395
5.0604 | -34.1496 | -9.0499
5.0665 | -34.15 | -9.0596
5.0726 | -34.1502 | -9.0701
5.0787 | -34.1506 | -9.0805
5.0848 | -34.1509 | -9.0906
5.091 | -34.1512 | -9.1006
5.0971 | -34.1515 | -9.1108
5.1032 | -34.1518 | -9.1212
5.1093 | -34.1522 | -9.1313
5.1154 | -34.1524 | -9.1413
5.1215 | -34.1526 | -9.1514
5.1276 | -34.153 | -9.1613
5.1337 | -34.1534 | -9.1711
5.1398 | -34.1537 | -9.181
5.1459 | -34.154 | -9.1915
5.152 | -34.1543 | -9.2026
5.1581 | -34.1545 | -9.2134
5.1642 | -34.1548 | -9.2234
5.1703 | -34.1552 | -9.2335
5.1764 | -34.1555 | -9.2435
5.1825 | -34.1558 | -9.2539
5.1886 | -34.1561 | -9.2638
5.1947 | -34.1563 | -9.2746
5.2008 | -34.1568 | -9.2846
5.2069 | -34.1571 | -9.2958
5.213 | -34.1574 | -9.3055
5.2191 | -34.1578 | -9.3159
5.2252 | -34.1581 | -9.3263
5.2314 | -34.1583 | -9.3371
5.2375 | -34.1587 | -9.3471
5.2436 | -34.159 | -9.3573
5.2497 | -34.1594 | -9.3673
5.2558 | -34.1597 | -9.3783
5.2619 | -34.16 | -9.3893
5.268 | -34.1604 | -9.3998
5.2741 | -34.1607 | -9.4111
5.2802 | -34.161 | -9.422
5.2863 | -34.1614 | -9.4322
5.2924 | -34.1617 | -9.4435
5.2985 | -34.1621 | -9.4549
5.3046 | -34.1626 | -9.4662
5.3107 | -34.1629 | -9.4771
5.3168 | -34.1634 | -9.4876
5.3229 | -34.1637 | -9.4976
5.329 | -34.1641 | -9.5076
5.3351 | -34.1645 | -9.5178
5.3412 | -34.1649 | -9.5281
5.3473 | -34.1653 | -9.539
5.3534 | -34.1657 | -9.55
5.3595 | -34.166 | -9.5608
5.3656 | -34.1663 | -9.5718
5.3717 | -34.1666 | -9.5824
5.3779 | -34.167 | -9.5934
5.384 | -34.1673 | -9.6042
5.3901 | -34.1677 | -9.615
5.3962 | -34.1679 | -9.6258
5.4023 | -34.1683 | -9.6367
5.4084 | -34.1687 | -9.6471
5.4145 | -34.169 | -9.6582
5.4206 | -34.1693 | -9.6695
5.4267 | -34.1697 | -9.681
5.4328 | -34.1701 | -9.6912
5.4389 | -34.1704 | -9.7015
5.445 | -34.1707 | -9.7129
5.4511 | -34.1711 | -9.7233
5.4572 | -34.1715 | -9.7346
5.4633 | -34.1718 | -9.745
5.4694 | -34.1721 | -9.7556
5.4755 | -34.1725 | -9.7665
5.4816 | -34.1728 | -9.7772
5.4877 | -34.1731 | -9.7879
5.4938 | -34.1735 | -9.7984
5.4999 | -34.1738 | -9.8091
5.506 | -34.1742 | -9.82
5.5121 | -34.1745 | -9.8311
5.5183 | -34.1748 | -9.8412
5.5244 | -34.1751 | -9.8521
5.5305 | -34.1755 | -9.8631
5.5366 | -34.1758 | -9.8735
5.5427 | -34.1761 | -9.8849
5.5488 | -34.1764 | -9.8959
5.5549 | -34.1767 | -9.9067
5.561 | -34.177 | -9.918
5.5671 | -34.1775 | -9.9285
5.5732 | -34.1777 | -9.9395
5.5793 | -34.1781 | -9.9511
5.5854 | -34.1785 | -9.962
5.5915 | -34.1787 | -9.9731
5.5976 | -34.179 | -9.9841
5.6037 | -34.1794 | -9.9949
5.6098 | -34.1797 | -10.0055
5.6159 | -34.18 | -10.0164
5.622 | -34.1804 | -10.027
5.6281 | -34.1807 | -10.0379
5.6342 | -34.1811 | -10.0488
5.6403 | -34.1813 | -10.0591
5.6464 | -34.1817 | -10.07
5.6525 | -34.182 | -10.0805
5.6586 | -34.1824 | -10.0906
5.6648 | -34.1827 | -10.1019
5.6709 | -34.183 | -10.1131
5.677 | -34.1834 | -10.1246
5.6831 | -34.1837 | -10.135
5.6892 | -34.184 | -10.1458
5.6953 | -34.1843 | -10.1565
5.7014 | -34.1847 | -10.1679
5.7075 | -34.1849 | -10.1789
5.7136 | -34.1853 | -10.1895
5.7197 | -34.1856 | -10.2001
5.7258 | -34.1858 | -10.2107
5.7319 | -34.1862 | -10.2215
5.738 | -34.1865 | -10.2323
5.7441 | -34.1869 | -10.243
5.7502 | -34.1872 | -10.2534
5.7563 | -34.1876 | -10.2639
5.7624 | -34.1879 | -10.2754
5.7685 | -34.1881 | -10.286
5.7746 | -34.1883 | -10.2966
5.7807 | -34.1887 | -10.3071
5.7868 | -34.1889 | -10.3178
5.7929 | -34.1892 | -10.3284
5.799 | -34.1896 | -10.339
5.8052 | -34.1898 | -10.3498
5.8113 | -34.1901 | -10.3606
5.8174 | -34.1904 | -10.3714
5.8235 | -34.1907 | -10.382
5.8296 | -34.191 | -10.3933
5.8357 | -34.1912 | -10.404
5.8418 | -34.1915 | -10.4149
5.8479 | -34.1917 | -10.4259
5.854 | -34.192 | -10.4369
5.8601 | -34.1923 | -10.4475
5.8662 | -34.1926 | -10.4585
5.8723 | -34.1929 | -10.4686
5.8784 | -34.1932 | -10.4793
5.8845 | -34.1935 | -10.4901
5.8906 | -34.1938 | -10.5007
5.8967 | -34.1941 | -10.511
5.9028 | -34.1944 | -10.5218
5.9089 | -34.1947 | -10.5319
5.915 | -34.195 | -10.5425
5.9211 | -34.1953 | -10.5532
5.9272 | -34.1956 | -10.5638
5.9333 | -34.1959 | -10.5738
5.9394 | -34.1962 | -10.5842
5.9455 | -34.1965 | -10.5945
5.9517 | -34.1969 | -10.6051
5.9578 | -34.1972 | -10.616
5.9639 | -34.1976 | -10.6263
5.97 | -34.1979 | -10.637
5.9761 | -34.1982 | -10.6473
5.9822 | -34.1985 | -10.6581
5.9883 | -34.1989 | -10.6684
5.9944 | -34.1992 | -10.6791
6.0005 | -34.1996 | -10.69
6.0066 | -34.1998 | -10.7002
6.0127 | -34.2001 | -10.71
6.0188 | -34.2004 | -10.721
6.0249 | -34.2007 | -10.7315
6.031 | -34.201 | -10.7418
6.0371 | -34.2014 | -10.7526
6.0432 | -34.2018 | -10.7631
6.0493 | -34.2021 | -10.7735
6.0554 | -34.2025 | -10.7839
6.0615 | -34.2028 | -10.795
6.0676 | -34.2031 | -10.8061
6.0737 | -34.2034 | -10.8168
6.0798 | -34.2037 | -10.8272
6.0859 | -34.204 | -10.8377
6.0921 | -34.2044 | -10.848
6.0982 | -34.2047 | -10.8588
6.1043 | -34.2051 | -10.8691
6.1104 | -34.2055 | -10.8802
6.1165 | -34.2058 | -10.8899
6.1226 | -34.2061 | -10.8999
6.1287 | -34.2065 | -10.9105
6.1348 | -34.2069 | -10.921
6.1409 | -34.2072 | -10.9314
6.147 | -34.2076 | -10.9417
6.1531 | -34.2079 | -10.9522
6.1592 | -34.2082 | -10.9626
6.1653 | -34.2085 | -10.9729
6.1714 | -34.2089 | -10.9834
6.1775 | -34.2092 | -10.9942
6.1836 | -34.2095 | -11.0044
6.1897 | -34.2099 | -11.015
6.1958 | -34.2101 | -11.0256
6.2019 | -34.2105 | -11.0364
6.208 | -34.2109 | -11.0472
6.2141 | -34.2112 | -11.0579
6.2202 | -34.2116 | -11.0688
6.2263 | -34.212 | -11.0792
6.2325 | -34.2123 | -11.0891
6.2386 | -34.2127 | -11.0992
6.2447 | -34.213 | -11.11
6.2508 | -34.2133 | -11.1203
6.2569 | -34.2137 | -11.1311
6.263 | -34.2141 | -11.1412
6.2691 | -34.2144 | -11.1515
6.2752 | -34.2148 | -11.1619
6.2813 | -34.2151 | -11.1723
6.2874 | -34.2154 | -11.1827
6.2935 | -34.2158 | -11.1929
6.2996 | -34.2162 | -11.2035
6.3057 | -34.2165 | -11.2138
6.3118 | -34.2169 | -11.2242
6.3179 | -34.2172 | -11.2346
6.324 | -34.2176 | -11.2453
6.3301 | -34.218 | -11.2555
6.3362 | -34.2183 | -11.266
6.3423 | -34.2188 | -11.2765
6.3484 | -34.2192 | -11.2871
6.3545 | -34.2195 | -11.2969
6.3606 | -34.2199 | -11.307
6.3667 | -34.2203 | -11.3172
6.3728 | -34.2207 | -11.3276
6.379 | -34.221 | -11.338
6.3851 | -34.2214 | -11.3484
6.3912 | -34.2218 | -11.3586
6.3973 | -34.2222 | -11.3687
6.4034 | -34.2225 | -11.379
6.4095 | -34.2228 | -11.3896
6.4156 | -34.2232 | -11.3997
6.4217 | -34.2235 | -11.41
6.4278 | -34.2239 | -11.4203
6.4339 | -34.2243 | -11.4306
6.44 | -34.2247 | -11.4406
6.4461 | -34.225 | -11.4514
6.4522 | -34.2254 | -11.4612
6.4583 | -34.2257 | -11.4716
6.4644 | -34.226 | -11.4823
6.4705 | -34.2264 | -11.4928
6.4766 | -34.2268 | -11.5036
6.4827 | -34.2272 | -11.5144
6.4888 | -34.2275 | -11.5245
6.4949 | -34.2279 | -11.5344
6.501 | -34.2282 | -11.5451
6.5071 | -34.2286 | -11.5556
6.5132 | -34.229 | -11.5658
6.5194 | -34.2294 | -11.5761
6.5255 | -34.2298 | -11.586
6.5316 | -34.2301 | -11.5965
6.5377 | -34.2304 | -11.6072
6.5438 | -34.2308 | -11.6174
6.5499 | -34.2312 | -11.6275
6.556 | -34.2316 | -11.6377
6.5621 | -34.2319 | -11.6482
6.5682 | -34.2322 | -11.6584
6.5743 | -34.2326 | -11.6695
6.5804 | -34.2329 | -11.6798
6.5865 | -34.2332 | -11.6904
6.5926 | -34.2336 | -11.7007
6.5987 | -34.234 | -11.7107
6.6048 | -34.2343 | -11.7217
6.6109 | -34.2347 | -11.732
6.617 | -34.2351 | -11.7432
6.6231 | -34.2353 | -11.7538
6.6292 | -34.2358 | -11.764
6.6353 | -34.2361 | -11.7754
6.6414 | -34.2366 | -11.786
6.6475 | -34.2369 | -11.7966
6.6536 | -34.2373 | -11.8066
6.6597 | -34.2377 | -11.8167
6.6659 | -34.2381 | -11.8269
6.672 | -34.2385 | -11.8373
6.6781 | -34.2388 | -11.8475
6.6842 | -34.2392 | -11.8581
6.6903 | -34.2396 | -11.8683
6.6964 | -34.24 | -11.8788
6.7025 | -34.2403 | -11.8891
6.7086 | -34.2407 | -11.8992
6.7147 | -34.241 | -11.9098
6.7208 | -34.2414 | -11.9207
6.7269 | -34.2417 | -11.9306
6.733 | -34.242 | -11.9402
6.7391 | -34.2425 | -11.9508
6.7452 | -34.2429 | -11.9606
6.7513 | -34.2433 | -11.9706
6.7574 | -34.2437 | -11.9812
6.7635 | -34.2441 | -11.9911
6.7696 | -34.2445 | -12.0015
6.7757 | -34.2449 | -12.0121
6.7818 | -34.2453 | -12.0225
6.7879 | -34.2456 | -12.033
6.794 | -34.246 | -12.0431
6.8001 | -34.2463 | -12.0532
6.8063 | -34.2467 | -12.0638
6.8124 | -34.2471 | -12.0742
6.8185 | -34.2474 | -12.0845
6.8246 | -34.2478 | -12.0948
6.8307 | -34.2481 | -12.105
6.8368 | -34.2485 | -12.1157
6.8429 | -34.2489 | -12.1259
6.849 | -34.2493 | -12.1362
6.8551 | -34.2497 | -12.1466
6.8612 | -34.2501 | -12.1568
6.8673 | -34.2504 | -12.1669
6.8734 | -34.2508 | -12.1779
6.8795 | -34.2512 | -12.1885
6.8856 | -34.2516 | -12.1983
6.8917 | -34.2519 | -12.2081
6.8978 | -34.2523 | -12.2183
6.9039 | -34.2526 | -12.2284
6.91 | -34.253 | -12.2389
6.9161 | -34.2534 | -12.249
6.9222 | -34.2537 | -12.2593
6.9283 | -34.2541 | -12.2692
6.9344 | -34.2544 | -12.2793
6.9405 | -34.2548 | -12.2894
6.9466 | -34.2552 | -12.2999
6.9528 | -34.2555 | -12.3102
6.9589 | -34.2559 | -12.3204
6.965 | -34.2563 | -12.3307
6.9711 | -34.2567 | -12.341
6.9772 | -34.2571 | -12.351
6.9833 | -34.2574 | -12.3614
6.9894 | -34.2577 | -12.3721
6.9955 | -34.2581 | -12.3823
7.0016 | -34.2585 | -12.393
7.0077 | -34.2589 | -12.4035
7.0138 | -34.2592 | -12.4141
7.0199 | -34.2596 | -12.4241
7.026 | -34.2598 | -12.4343
7.0321 | -34.2602 | -12.4445
7.0382 | -34.2605 | -12.4545
7.0443 | -34.261 | -12.4651
7.0504 | -34.2613 | -12.4756
7.0565 | -34.2616 | -12.4861
7.0626 | -34.262 | -12.4961
7.0687 | -34.2623 | -12.506
7.0748 | -34.2627 | -12.516
7.0809 | -34.263 | -12.5264
7.087 | -34.2634 | -12.5367
7.0932 | -34.2637 | -12.5469
7.0993 | -34.264 | -12.5572
7.1054 | -34.2644 | -12.5672
7.1115 | -34.2647 | -12.5773
7.1176 | -34.265 | -12.5873
7.1237 | -34.2653 | -12.5975
7.1298 | -34.2656 | -12.6073
7.1359 | -34.2659 | -12.6172
7.142 | -34.2662 | -12.6274
7.1481 | -34.2665 | -12.6373
7.1542 | -34.2668 | -12.6474
7.1603 | -34.2672 | -12.6574
7.1664 | -34.2676 | -12.6676
7.1725 | -34.2679 | -12.6776
7.1786 | -34.2683 | -12.6879
7.1847 | -34.2686 | -12.6979
7.1908 | -34.2689 | -12.7081
7.1969 | -34.2693 | -12.7182
7.203 | -34.2696 | -12.7284
7.2091 | -34.2699 | -12.7382
7.2152 | -34.2703 | -12.7484
7.2213 | -34.2706 | -12.7583
7.2274 | -34.2709 | -12.7684
7.2335 | -34.2713 | -12.7785
7.2397 | -34.2716 | -12.7886
7.2458 | -34.2719 | -12.7982
7.2519 | -34.2722 | -12.8083
7.258 | -34.2726 | -12.8185
7.2641 | -34.2729 | -12.8285
7.2702 | -34.2733 | -12.8389
7.2763 | -34.2736 | -12.8487
7.2824 | -34.274 | -12.8587
7.2885 | -34.2744 | -12.8687
7.2946 | -34.2747 | -12.8789
7.3007 | -34.275 | -12.889
7.3068 | -34.2753 | -12.8993
7.3129 | -34.2756 | -12.9091
7.319 | -34.276 | -12.9193
7.3251 | -34.2763 | -12.929
7.3312 | -34.2767 | -12.9387
7.3373 | -34.2771 | -12.9484
7.3434 | -34.2775 | -12.9582
7.3495 | -34.2778 | -12.9677
7.3556 | -34.2781 | -12.9782
7.3617 | -34.2784 | -12.9886
7.3678 | -34.2788 | -12.9987
7.3739 | -34.2791 | -13.0083
7.3801 | -34.2794 | -13.0183
7.3862 | -34.2797 | -13.0282
7.3923 | -34.2801 | -13.0381
7.3984 | -34.2804 | -13.048
7.4045 | -34.2807 | -13.0579
7.4106 | -34.2811 | -13.0677
7.4167 | -34.2814 | -13.0779
7.4228 | -34.2818 | -13.0878
7.4289 | -34.2821 | -13.0974
7.435 | -34.2825 | -13.1071
7.4411 | -34.2828 | -13.1168
7.4472 | -34.2831 | -13.1268
7.4533 | -34.2835 | -13.1367
7.4594 | -34.2838 | -13.147
7.4655 | -34.2842 | -13.1569
7.4716 | -34.2845 | -13.167
7.4777 | -34.2848 | -13.1764
7.4838 | -34.2852 | -13.1865
7.4899 | -34.2856 | -13.1964
7.496 | -34.2859 | -13.2064
7.5021 | -34.2863 | -13.2162
7.5082 | -34.2866 | -13.2262
7.5143 | -34.2869 | -13.236
7.5204 | -34.2872 | -13.2456
7.5266 | -34.2875 | -13.2555
7.5327 | -34.2879 | -13.2653
7.5388 | -34.2882 | -13.2751
7.5449 | -34.2885 | -13.2852
7.551 | -34.2888 | -13.2954
7.5571 | -34.2891 | -13.3052
7.5632 | -34.2895 | -13.315
7.5693 | -34.2898 | -13.325
7.5754 | -34.2902 | -13.3349
7.5815 | -34.2906 | -13.3444
7.5876 | -34.2908 | -13.3543
7.5937 | -34.2912 | -13.364
7.5998 | -34.2915 | -13.374
7.6059 | -34.2918 | -13.3838
7.612 | -34.2921 | -13.3936
7.6181 | -34.2925 | -13.4034
7.6242 | -34.2928 | -13.4134
7.6303 | -34.2932 | -13.423
7.6364 | -34.2935 | -13.4331
7.6425 | -34.2938 | -13.4433
7.6486 | -34.2942 | -13.4532
7.6547 | -34.2945 | -13.4632
7.6608 | -34.2949 | -13.473
7.667 | -34.2952 | -13.4823
7.6731 | -34.2955 | -13.4923
7.6792 | -34.2958 | -13.5024
7.6853 | -34.2961 | -13.5121
7.6914 | -34.2963 | -13.5221
7.6975 | -34.2967 | -13.5321
7.7036 | -34.297 | -13.5421
7.7097 | -34.2973 | -13.5519
7.7158 | -34.2977 | -13.5616
7.7219 | -34.2979 | -13.5713
7.728 | -34.2982 | -13.5813
7.7341 | -34.2985 | -13.5912
7.7402 | -34.2989 | -13.601
7.7463 | -34.2992 | -13.6109
7.7524 | -34.2995 | -13.6211
7.7585 | -34.2998 | -13.6311
7.7646 | -34.3001 | -13.6406
7.7707 | -34.3004 | -13.6503
7.7768 | -34.3006 | -13.66
7.7829 | -34.3009 | -13.67
7.789 | -34.3012 | -13.6796
7.7951 | -34.3015 | -13.6898
7.8012 | -34.3018 | -13.7
7.8073 | -34.3021 | -13.7097
7.8135 | -34.3024 | -13.7197
7.8196 | -34.3028 | -13.7295
7.8257 | -34.3031 | -13.7397
7.8318 | -34.3034 | -13.75
7.8379 | -34.3037 | -13.7594
7.844 | -34.304 | -13.7688
7.8501 | -34.3042 | -13.7789
7.8562 | -34.3045 | -13.7894
7.8623 | -34.3049 | -13.7993
7.8684 | -34.3051 | -13.8089
7.8745 | -34.3054 | -13.8186
7.8806 | -34.3057 | -13.828
7.8867 | -34.306 | -13.8376
7.8928 | -34.3064 | -13.8473
7.8989 | -34.3067 | -13.8567
7.905 | -34.3071 | -13.867
7.9111 | -34.3074 | -13.8769
7.9172 | -34.3077 | -13.8865
7.9233 | -34.3079 | -13.8963
7.9294 | -34.3083 | -13.9063
7.9355 | -34.3086 | -13.9157
7.9416 | -34.3088 | -13.9254
7.9477 | -34.3091 | -13.9348
7.9539 | -34.3094 | -13.9442
7.96 | -34.3097 | -13.9543
7.9661 | -34.31 | -13.9642
7.9722 | -34.3104 | -13.9741
7.9783 | -34.3107 | -13.984
7.9844 | -34.311 | -13.9942
7.9905 | -34.3113 | -14.0039
7.9966 | -34.3115 | -14.0137
8.0027 | -34.3119 | -14.0238
8.0088 | -34.3122 | -14.0333
8.0149 | -34.3124 | -14.0428
8.021 | -34.3127 | -14.0527
8.0271 | -34.3131 | -14.0629
8.0332 | -34.3133 | -14.0727
8.0393 | -34.3137 | -14.0827
8.0454 | -34.3139 | -14.0927
8.0515 | -34.3143 | -14.1025
8.0576 | -34.3145 | -14.1126
8.0637 | -34.3148 | -14.1219
8.0698 | -34.3151 | -14.1313
8.0759 | -34.3154 | -14.1408
8.082 | -34.3158 | -14.1508
8.0881 | -34.316 | -14.1605
8.0942 | -34.3163 | -14.1708
8.1004 | -34.3166 | -14.1809
8.1065 | -34.3169 | -14.1906
8.1126 | -34.3173 | -14.2004
8.1187 | -34.3176 | -14.2103
8.1248 | -34.3179 | -14.2203
8.1309 | -34.3182 | -14.2301
8.137 | -34.3185 | -14.2397
8.1431 | -34.3188 | -14.2497
8.1492 | -34.3191 | -14.2596
8.1553 | -34.3194 | -14.2695
8.1614 | -34.3197 | -14.2796
8.1675 | -34.32 | -14.2892
8.1736 | -34.3203 | -14.2993
8.1797 | -34.3205 | -14.3093
8.1858 | -34.3208 | -14.3191
8.1919 | -34.3211 | -14.3289
8.198 | -34.3214 | -14.3383
8.2041 | -34.3217 | -14.3485
8.2102 | -34.3221 | -14.3583
8.2163 | -34.3224 | -14.3681
8.2224 | -34.3227 | -14.3781
8.2285 | -34.323 | -14.3878
8.2346 | -34.3233 | -14.3975
8.2408 | -34.3236 | -14.4074
8.2469 | -34.3239 | -14.4171
8.253 | -34.3242 | -14.4269
8.2591 | -34.3245 | -14.4364
8.2652 | -34.3249 | -14.4461
8.2713 | -34.3252 | -14.456
8.2774 | -34.3255 | -14.4658
8.2835 | -34.3258 | -14.4761
8.2896 | -34.3261 | -14.4854
8.2957 | -34.3264 | -14.4952
8.3018 | -34.3267 | -14.505
8.3079 | -34.3271 | -14.5148
8.314 | -34.3274 | -14.5245
8.3201 | -34.3277 | -14.5342
8.3262 | -34.328 | -14.544
8.3323 | -34.3283 | -14.5538
8.3384 | -34.3286 | -14.564
8.3445 | -34.3289 | -14.5733
8.3506 | -34.3293 | -14.5834
8.3567 | -34.3296 | -14.5934
8.3628 | -34.3298 | -14.6032
8.3689 | -34.3301 | -14.613
8.375 | -34.3304 | -14.6229
8.3812 | -34.3308 | -14.6326
8.3873 | -34.3311 | -14.6425
8.3934 | -34.3314 | -14.6523
8.3995 | -34.3317 | -14.6622
8.4056 | -34.332 | -14.6721
8.4117 | -34.3323 | -14.6822
8.4178 | -34.3326 | -14.6916
8.4239 | -34.3329 | -14.7012
8.43 | -34.3333 | -14.711
8.4361 | -34.3336 | -14.7211
8.4422 | -34.3339 | -14.7309
8.4483 | -34.3341 | -14.7407
8.4544 | -34.3345 | -14.7508
8.4605 | -34.3348 | -14.7605
8.4666 | -34.335 | -14.7705
8.4727 | -34.3353 | -14.7802
8.4788 | -34.3356 | -14.7899
8.4849 | -34.3359 | -14.8
8.491 | -34.3362 | -14.8097
8.4971 | -34.3365 | -14.8196
8.5032 | -34.3369 | -14.8294
8.5093 | -34.3372 | -14.8393
8.5154 | -34.3375 | -14.8489
8.5215 | -34.3378 | -14.8588
8.5277 | -34.3381 | -14.8687
8.5338 | -34.3384 | -14.8782
8.5399 | -34.3387 | -14.8883
8.546 | -34.339 | -14.8978
8.5521 | -34.3393 | -14.9078
8.5582 | -34.3396 | -14.9177
8.5643 | -34.3399 | -14.9274
8.5704 | -34.3402 | -14.9368
8.5765 | -34.3405 | -14.9466
8.5826 | -34.3409 | -14.9564
8.5887 | -34.3412 | -14.9661
8.5948 | -34.3415 | -14.9755
8.6009 | -34.3418 | -14.9853
8.607 | -34.3421 | -14.9954
8.6131 | -34.3424 | -15.0054
8.6192 | -34.3427 | -15.0149
8.6253 | -34.3431 | -15.0248
8.6314 | -34.3434 | -15.0346
8.6375 | -34.3437 | -15.0446
8.6436 | -34.3441 | -15.0544
8.6497 | -34.3444 | -15.0643
8.6558 | -34.3447 | -15.0738
8.6619 | -34.345 | -15.0839
8.6681 | -34.3452 | -15.0936
8.6742 | -34.3456 | -15.1036
8.6803 | -34.3459 | -15.1137
8.6864 | -34.3463 | -15.1238
8.6925 | -34.3466 | -15.1336
8.6986 | -34.3469 | -15.1437
8.7047 | -34.3472 | -15.1536
8.7108 | -34.3475 | -15.1635
8.7169 | -34.3478 | -15.1731
8.723 | -34.3482 | -15.1829
8.7291 | -34.3485 | -15.1926
8.7352 | -34.3489 | -15.2027
8.7413 | -34.3492 | -15.2124
8.7474 | -34.3495 | -15.222
8.7535 | -34.3499 | -15.232
8.7596 | -34.3502 | -15.2424
8.7657 | -34.3505 | -15.2522
8.7718 | -34.3508 | -15.2617
8.7779 | -34.3512 | -15.272
8.784 | -34.3514 | -15.2817
8.7901 | -34.3517 | -15.2919
8.7962 | -34.3521 | -15.3019
8.8023 | -34.3524 | -15.3113
8.8084 | -34.3527 | -15.3213
8.8146 | -34.353 | -15.3311
8.8207 | -34.3534 | -15.3409
8.8268 | -34.3537 | -15.3507
8.8329 | -34.3541 | -15.3606
8.839 | -34.3544 | -15.3705
8.8451 | -34.3547 | -15.3801
8.8512 | -34.3551 | -15.3901
8.8573 | -34.3554 | -15.3997
8.8634 | -34.3558 | -15.4095
8.8695 | -34.3561 | -15.4194
8.8756 | -34.3564 | -15.4291
8.8817 | -34.3568 | -15.4391
8.8878 | -34.357 | -15.4492
8.8939 | -34.3573 | -15.4587
8.9 | -34.3577 | -15.4683
8.9061 | -34.358 | -15.4781
8.9122 | -34.3584 | -15.4879
8.9183 | -34.3587 | -15.4976
8.9244 | -34.359 | -15.5072
8.9305 | -34.3594 | -15.517
8.9366 | -34.3598 | -15.5268
8.9427 | -34.3601 | -15.5367
8.9488 | -34.3605 | -15.5465
8.955 | -34.3608 | -15.5563
8.9611 | -34.3611 | -15.5662
8.9672 | -34.3615 | -15.5764
8.9733 | -34.3618 | -15.5861
8.9794 | -34.3622 | -15.596
8.9855 | -34.3625 | -15.6058
8.9916 | -34.3629 | -15.6155
8.9977 | -34.3632 | -15.6252
9.0038 | -34.3636 | -15.6352
9.0099 | -34.364 | -15.6452
9.016 | -34.3643 | -15.655
9.0221 | -34.3647 | -15.6651
9.0282 | -34.365 | -15.675
9.0343 | -34.3654 | -15.6851
9.0404 | -34.3658 | -15.6949
9.0465 | -34.3662 | -15.705
9.0526 | -34.3665 | -15.7148
9.0587 | -34.3669 | -15.7246
9.0648 | -34.3673 | -15.7344
9.0709 | -34.3676 | -15.7439
9.077 | -34.368 | -15.7543
9.0831 | -34.3683 | -15.7642
9.0892 | -34.3686 | -15.774
9.0953 | -34.3689 | -15.7839
9.1015 | -34.3693 | -15.7937
9.1076 | -34.3697 | -15.8037
9.1137 | -34.3701 | -15.8135
9.1198 | -34.3704 | -15.823
9.1259 | -34.3708 | -15.8329
9.132 | -34.3712 | -15.8429
9.1381 | -34.3716 | -15.8526
9.1442 | -34.372 | -15.862
9.1503 | -34.3723 | -15.872
9.1564 | -34.3727 | -15.8818
9.1625 | -34.373 | -15.8918
9.1686 | -34.3734 | -15.9016
9.1747 | -34.3738 | -15.9116
9.1808 | -34.3742 | -15.9217
9.1869 | -34.3745 | -15.9318
9.193 | -34.3749 | -15.9418
9.1991 | -34.3752 | -15.9515
9.2052 | -34.3756 | -15.9612
9.2113 | -34.376 | -15.9713
9.2174 | -34.3764 | -15.9817
9.2235 | -34.3768 | -15.9913
9.2296 | -34.3772 | -16.0013
9.2357 | -34.3775 | -16.0111
9.2419 | -34.3779 | -16.0209
9.248 | -34.3783 | -16.031
9.2541 | -34.3787 | -16.041
9.2602 | -34.379 | -16.051
9.2663 | -34.3794 | -16.0606
9.2724 | -34.3798 | -16.0705
9.2785 | -34.3801 | -16.0803
9.2846 | -34.3805 | -16.0902
9.2907 | -34.3809 | -16.1
9.2968 | -34.3812 | -16.11
9.3029 | -34.3816 | -16.12
9.309 | -34.382 | -16.1298
9.3151 | -34.3823 | -16.1394
9.3212 | -34.3828 | -16.1494
9.3273 | -34.3832 | -16.159
9.3334 | -34.3836 | -16.1686
9.3395 | -34.384 | -16.1783
9.3456 | -34.3844 | -16.188
9.3517 | -34.3848 | -16.1974
9.3578 | -34.3852 | -16.207
9.3639 | -34.3856 | -16.2171
9.37 | -34.386 | -16.2271
9.3761 | -34.3864 | -16.2369
9.3822 | -34.3868 | -16.2466
9.3884 | -34.3872 | -16.2564
9.3945 | -34.3876 | -16.2663
9.4006 | -34.3879 | -16.2762
9.4067 | -34.3883 | -16.2862
9.4128 | -34.3887 | -16.2959
9.4189 | -34.3891 | -16.3059
9.425 | -34.3895 | -16.3155
9.4311 | -34.3899 | -16.3254
9.4372 | -34.3903 | -16.3353
9.4433 | -34.3907 | -16.3453
9.4494 | -34.3911 | -16.3551
9.4555 | -34.3915 | -16.3649
9.4616 | -34.3919 | -16.3746
9.4677 | -34.3923 | -16.3841
9.4738 | -34.3927 | -16.3943
9.4799 | -34.3931 | -16.4041
9.486 | -34.3934 | -16.4142
9.4921 | -34.3938 | -16.4244
9.4982 | -34.3943 | -16.4337
9.5043 | -34.3947 | -16.4432
9.5104 | -34.3951 | -16.4529
9.5165 | -34.3955 | -16.4627
9.5226 | -34.3959 | -16.4726
9.5288 | -34.3963 | -16.4826
9.5349 | -34.3967 | -16.4924
9.541 | -34.3971 | -16.5022
9.5471 | -34.3975 | -16.5121
9.5532 | -34.3979 | -16.5219
9.5593 | -34.3983 | -16.5316
9.5654 | -34.3987 | -16.5415
9.5715 | -34.3991 | -16.5516
9.5776 | -34.3995 | -16.5618
9.5837 | -34.3999 | -16.5714
9.5898 | -34.4002 | -16.5813
9.5959 | -34.4006 | -16.591
9.602 | -34.401 | -16.6009
9.6081 | -34.4014 | -16.6108
9.6142 | -34.4018 | -16.6205
9.6203 | -34.4022 | -16.6304
9.6264 | -34.4025 | -16.6399
9.6325 | -34.4029 | -16.6498
9.6386 | -34.4033 | -16.6595
9.6447 | -34.4037 | -16.6694
9.6508 | -34.4041 | -16.6792
9.6569 | -34.4045 | -16.689
9.663 | -34.4049 | -16.699
9.6691 | -34.4054 | -16.7088
9.6753 | -34.4058 | -16.7186
9.6814 | -34.4062 | -16.7281
9.6875 | -34.4067 | -16.7377
9.6936 | -34.4071 | -16.7473
9.6997 | -34.4075 | -16.7573
9.7058 | -34.4079 | -16.7671
9.7119 | -34.4083 | -16.7768
9.718 | -34.4087 | -16.7869
9.7241 | -34.4091 | -16.7967
9.7302 | -34.4094 | -16.8065
9.7363 | -34.4099 | -16.8161
9.7424 | -34.4103 | -16.8256
9.7485 | -34.4107 | -16.8352
9.7546 | -34.4111 | -16.8447
9.7607 | -34.4115 | -16.8546
9.7668 | -34.4119 | -16.8647
9.7729 | -34.4124 | -16.8741
9.779 | -34.4128 | -16.8839
9.7851 | -34.4132 | -16.8939
9.7912 | -34.4136 | -16.9041
9.7973 | -34.414 | -16.9141
9.8034 | -34.4144 | -16.9241
9.8095 | -34.4148 | -16.934
9.8157 | -34.4153 | -16.9437
9.8218 | -34.4157 | -16.9531
9.8279 | -34.4161 | -16.963
9.834 | -34.4165 | -16.9728
9.8401 | -34.417 | -16.9824
9.8462 | -34.4174 | -16.9919
9.8523 | -34.4179 | -17.0015
9.8584 | -34.4183 | -17.0115
9.8645 | -34.4187 | -17.0213
9.8706 | -34.4191 | -17.031
9.8767 | -34.4194 | -17.0404
9.8828 | -34.4199 | -17.0497
9.8889 | -34.4203 | -17.0597
9.895 | -34.4207 | -17.069
9.9011 | -34.4211 | -17.0783
9.9072 | -34.4215 | -17.0879
9.9133 | -34.4219 | -17.0973
9.9194 | -34.4223 | -17.1068
9.9255 | -34.4227 | -17.1166
9.9316 | -34.4232 | -17.1262
9.9377 | -34.4236 | -17.1356
9.9438 | -34.424 | -17.1453
9.9499 | -34.4244 | -17.1553
9.956 | -34.4249 | -17.1646
9.9622 | -34.4253 | -17.1743
9.9683 | -34.4257 | -17.1842
9.9744 | -34.4261 | -17.1939
9.9805 | -34.4266 | -17.2036
9.9866 | -34.427 | -17.2129
9.9927 | -34.4274 | -17.2227
9.9988 | -34.4279 | -17.2322
10.0049 | -34.4282 | -17.242
10.011 | -34.4287 | -17.2516
10.0171 | -34.4291 | -17.2614
10.0232 | -34.4295 | -17.2707
10.0293 | -34.43 | -17.2801
10.0354 | -34.4305 | -17.2895
10.0415 | -34.4309 | -17.2991
10.0476 | -34.4313 | -17.309
10.0537 | -34.4317 | -17.3188
10.0598 | -34.4321 | -17.3286
10.0659 | -34.4325 | -17.3382
10.072 | -34.433 | -17.3477
10.0781 | -34.4334 | -17.3573
10.0842 | -34.4338 | -17.367
10.0903 | -34.4342 | -17.3769
10.0964 | -34.4347 | -17.3862
10.1026 | -34.4352 | -17.3959
10.1087 | -34.4356 | -17.4055
10.1148 | -34.436 | -17.4153
10.1209 | -34.4365 | -17.4247
10.127 | -34.4369 | -17.434
10.1331 | -34.4373 | -17.4434
10.1392 | -34.4377 | -17.4529
10.1453 | -34.4382 | -17.4626
10.1514 | -34.4386 | -17.4719
10.1575 | -34.439 | -17.4813
10.1636 | -34.4395 | -17.4902
10.1697 | -34.4399 | -17.4997
10.1758 | -34.4404 | -17.5092
10.1819 | -34.4408 | -17.5189
10.188 | -34.4412 | -17.5282
10.1941 | -34.4416 | -17.538
10.2002 | -34.442 | -17.5469
10.2063 | -34.4424 | -17.5566
10.2124 | -34.4428 | -17.5662
10.2185 | -34.4433 | -17.5757
10.2246 | -34.4437 | -17.5852
10.2307 | -34.4441 | -17.5949
10.2368 | -34.4446 | -17.6035
10.2429 | -34.445 | -17.6128
10.2491 | -34.4454 | -17.6221
10.2552 | -34.4458 | -17.632
10.2613 | -34.4463 | -17.6417
10.2674 | -34.4467 | -17.6513
10.2735 | -34.4471 | -17.6611
10.2796 | -34.4475 | -17.6706
10.2857 | -34.4479 | -17.6803
10.2918 | -34.4483 | -17.6899
10.2979 | -34.4488 | -17.6997
10.304 | -34.4493 | -17.7093
10.3101 | -34.4497 | -17.7188
10.3162 | -34.4501 | -17.7283
10.3223 | -34.4505 | -17.7378
10.3284 | -34.451 | -17.7472
10.3345 | -34.4514 | -17.7566
10.3406 | -34.4518 | -17.7661
10.3467 | -34.4522 | -17.7756
10.3528 | -34.4526 | -17.785
10.3589 | -34.453 | -17.7947
10.365 | -34.4535 | -17.8046
10.3711 | -34.4538 | -17.8143
10.3772 | -34.4542 | -17.8239
10.3833 | -34.4546 | -17.8333
10.3895 | -34.455 | -17.8426
10.3956 | -34.4555 | -17.8519
10.4017 | -34.4559 | -17.8613
10.4078 | -34.4563 | -17.8708
10.4139 | -34.4567 | -17.8802
10.42 | -34.4572 | -17.8896
10.4261 | -34.4576 | -17.899
10.4322 | -34.4581 | -17.9083
10.4383 | -34.4584 | -17.9178
10.4444 | -34.4589 | -17.9269
10.4505 | -34.4593 | -17.9365
10.4566 | -34.4597 | -17.9458
10.4627 | -34.4601 | -17.9552
10.4688 | -34.4606 | -17.9644
10.4749 | -34.461 | -17.9737
10.481 | -34.4614 | -17.9832
10.4871 | -34.4618 | -17.9927
10.4932 | -34.4623 | -18.0021
10.4993 | -34.4628 | -18.0115
10.5054 | -34.4632 | -18.021
10.5115 | -34.4636 | -18.0304
10.5176 | -34.4641 | -18.0395
10.5237 | -34.4645 | -18.0489
10.5298 | -34.4649 | -18.0581
10.536 | -34.4654 | -18.0676
10.5421 | -34.4658 | -18.0766
10.5482 | -34.4662 | -18.0864
10.5543 | -34.4667 | -18.0958
10.5604 | -34.4671 | -18.1049
10.5665 | -34.4676 | -18.1141
10.5726 | -34.468 | -18.1233
10.5787 | -34.4684 | -18.1329
10.5848 | -34.4689 | -18.1418
10.5909 | -34.4692 | -18.1513
10.597 | -34.4696 | -18.1607
10.6031 | -34.4701 | -18.1702
10.6092 | -34.4705 | -18.1799
10.6153 | -34.4709 | -18.1895
10.6214 | -34.4714 | -18.1988
10.6275 | -34.4718 | -18.208
10.6336 | -34.4723 | -18.2177
10.6397 | -34.4727 | -18.227
10.6458 | -34.473 | -18.2361
10.6519 | -34.4734 | -18.2453
10.658 | -34.4738 | -18.2545
10.6641 | -34.4743 | -18.2639
10.6702 | -34.4747 | -18.2733
10.6764 | -34.4752 | -18.2825
10.6825 | -34.4756 | -18.292
10.6886 | -34.4761 | -18.3014
10.6947 | -34.4765 | -18.3109
10.7008 | -34.4769 | -18.3201
10.7069 | -34.4774 | -18.3296
10.713 | -34.4778 | -18.3388
10.7191 | -34.4783 | -18.3483
10.7252 | -34.4787 | -18.3574
10.7313 | -34.4792 | -18.3668
10.7374 | -34.4796 | -18.3759
10.7435 | -34.48 | -18.3855
10.7496 | -34.4804 | -18.3946
10.7557 | -34.4808 | -18.4039
10.7618 | -34.4813 | -18.4132
10.7679 | -34.4817 | -18.4224
10.774 | -34.4821 | -18.4317
10.7801 | -34.4825 | -18.4414
10.7862 | -34.483 | -18.4506
10.7923 | -34.4834 | -18.46
10.7984 | -34.4839 | -18.4693
10.8045 | -34.4843 | -18.4787
10.8106 | -34.4847 | -18.4881
10.8168 | -34.4851 | -18.4973
10.8229 | -34.4856 | -18.5068
10.829 | -34.486 | -18.5162
10.8351 | -34.4864 | -18.5253
10.8412 | -34.4868 | -18.5345
10.8473 | -34.4872 | -18.5436
10.8534 | -34.4876 | -18.5531
10.8595 | -34.488 | -18.5621
10.8656 | -34.4885 | -18.5713
10.8717 | -34.4889 | -18.5802
10.8778 | -34.4894 | -18.5894
10.8839 | -34.4898 | -18.5985
10.89 | -34.4902 | -18.6076
10.8961 | -34.4907 | -18.617
10.9022 | -34.4911 | -18.626
10.9083 | -34.4915 | -18.6356
10.9144 | -34.492 | -18.6447
10.9205 | -34.4925 | -18.654
10.9266 | -34.493 | -18.6636
10.9327 | -34.4934 | -18.6729
10.9388 | -34.4938 | -18.6821
10.9449 | -34.4943 | -18.6915
10.951 | -34.4947 | -18.7007
10.9571 | -34.4951 | -18.7101
10.9633 | -34.4955 | -18.7196
10.9694 | -34.496 | -18.7292
10.9755 | -34.4964 | -18.7386
10.9816 | -34.4968 | -18.7479
10.9877 | -34.4973 | -18.7572
10.9938 | -34.4977 | -18.7667
10.9999 | -34.4981 | -18.7759
11.006 | -34.4986 | -18.7851
11.0121 | -34.499 | -18.7947
11.0182 | -34.4994 | -18.8037
11.0243 | -34.4998 | -18.813
11.0304 | -34.5002 | -18.8222
11.0365 | -34.5007 | -18.8315
11.0426 | -34.5011 | -18.8405
11.0487 | -34.5016 | -18.8498
11.0548 | -34.502 | -18.859
11.0609 | -34.5025 | -18.8686
11.067 | -34.5029 | -18.8781
11.0731 | -34.5033 | -18.8872
11.0792 | -34.5038 | -18.8965
11.0853 | -34.5042 | -18.9055
11.0914 | -34.5046 | -18.9147
11.0975 | -34.5051 | -18.9241
11.1037 | -34.5055 | -18.9332
11.1098 | -34.5059 | -18.9427
11.1159 | -34.5064 | -18.952
11.122 | -34.5068 | -18.9616
11.1281 | -34.5072 | -18.9705
11.1342 | -34.5077 | -18.9794
11.1403 | -34.5081 | -18.989
11.1464 | -34.5086 | -18.9982
11.1525 | -34.509 | -19.0075
11.1586 | -34.5095 | -19.0172
11.1647 | -34.5099 | -19.0263
11.1708 | -34.5103 | -19.0354
11.1769 | -34.5107 | -19.0441
11.183 | -34.5112 | -19.0534
11.1891 | -34.5117 | -19.063
11.1952 | -34.5121 | -19.0726
11.2013 | -34.5125 | -19.0819
11.2074 | -34.5129 | -19.0908
11.2135 | -34.5133 | -19.1002
11.2196 | -34.5137 | -19.1095
11.2257 | -34.5141 | -19.1186
11.2318 | -34.5144 | -19.1281
11.2379 | -34.5148 | -19.1372
11.244 | -34.5153 | -19.1465
11.2502 | -34.5157 | -19.1556
11.2563 | -34.5162 | -19.165
11.2624 | -34.5166 | -19.1742
11.2685 | -34.517 | -19.1833
11.2746 | -34.5174 | -19.1921
11.2807 | -34.5178 | -19.2019
11.2868 | -34.5182 | -19.2116
11.2929 | -34.5187 | -19.2205
11.299 | -34.5192 | -19.2301
11.3051 | -34.5196 | -19.2394
11.3112 | -34.52 | -19.2489
11.3173 | -34.5205 | -19.2583
11.3234 | -34.5208 | -19.2676
11.3295 | -34.5213 | -19.2767
11.3356 | -34.5218 | -19.2859
11.3417 | -34.5222 | -19.2955
11.3478 | -34.5226 | -19.3047
11.3539 | -34.523 | -19.3137
11.36 | -34.5234 | -19.3228
11.3661 | -34.5239 | -19.3321
11.3722 | -34.5243 | -19.3412
11.3783 | -34.5247 | -19.3506
11.3844 | -34.5251 | -19.3599
11.3906 | -34.5255 | -19.3695
11.3967 | -34.5259 | -19.3785
11.4028 | -34.5263 | -19.388
11.4089 | -34.5268 | -19.397
11.415 | -34.5271 | -19.4062
11.4211 | -34.5276 | -19.4151
11.4272 | -34.528 | -19.4243
11.4333 | -34.5284 | -19.4338
11.4394 | -34.5288 | -19.4433
11.4455 | -34.5292 | -19.4526
11.4516 | -34.5296 | -19.4617
11.4577 | -34.53 | -19.4712
11.4638 | -34.5305 | -19.4802
11.4699 | -34.5309 | -19.4897
11.476 | -34.5313 | -19.4994
11.4821 | -34.5318 | -19.5084
11.4882 | -34.5322 | -19.518
11.4943 | -34.5326 | -19.5271
11.5004 | -34.5331 | -19.5365
11.5065 | -34.5335 | -19.5458
11.5126 | -34.534 | -19.5546
11.5187 | -34.5344 | -19.5638
11.5248 | -34.5349 | -19.5731
11.5309 | -34.5353 | -19.5822
11.5371 | -34.5357 | -19.5913
11.5432 | -34.5361 | -19.6008
11.5493 | -34.5365 | -19.6102
11.5554 | -34.5369 | -19.6194
11.5615 | -34.5373 | -19.6289
11.5676 | -34.5378 | -19.6384
11.5737 | -34.5383 | -19.6472
11.5798 | -34.5387 | -19.6562
11.5859 | -34.5392 | -19.6654
11.592 | -34.5396 | -19.6747
11.5981 | -34.54 | -19.6836
11.6042 | -34.5404 | -19.6932
11.6103 | -34.5408 | -19.7024
11.6164 | -34.5412 | -19.7113
11.6225 | -34.5416 | -19.7207
11.6286 | -34.5421 | -19.7301
11.6347 | -34.5425 | -19.7396
11.6408 | -34.543 | -19.7487
11.6469 | -34.5434 | -19.7581
11.653 | -34.5439 | -19.7671
11.6591 | -34.5443 | -19.7762
11.6652 | -34.5447 | -19.7854
11.6713 | -34.5451 | -19.7947
11.6775 | -34.5456 | -19.804
11.6836 | -34.546 | -19.8132
11.6897 | -34.5465 | -19.8222
11.6958 | -34.5469 | -19.8318
11.7019 | -34.5474 | -19.8411
11.708 | -34.5478 | -19.8504
11.7141 | -34.5483 | -19.8597
11.7202 | -34.5487 | -19.8686
11.7263 | -34.5491 | -19.8779
11.7324 | -34.5495 | -19.8874
11.7385 | -34.5499 | -19.8968
11.7446 | -34.5503 | -19.9062
11.7507 | -34.5508 | -19.9155
11.7568 | -34.5512 | -19.9248
11.7629 | -34.5516 | -19.9341
11.769 | -34.5521 | -19.9433
11.7751 | -34.5525 | -19.9523
11.7812 | -34.553 | -19.9614
11.7873 | -34.5534 | -19.9709
11.7934 | -34.5539 | -19.9802
11.7995 | -34.5543 | -19.989
11.8056 | -34.5547 | -19.9985
11.8117 | -34.555 | -20.0077
11.8178 | -34.5555 | -20.0172
11.824 | -34.556 | -20.0262
11.8301 | -34.5565 | -20.035
11.8362 | -34.5569 | -20.0439
11.8423 | -34.5573 | -20.0532
11.8484 | -34.5577 | -20.0623
11.8545 | -34.5582 | -20.0718
11.8606 | -34.5586 | -20.0809
11.8667 | -34.559 | -20.0902
11.8728 | -34.5595 | -20.0993
11.8789 | -34.5599 | -20.1089
11.885 | -34.5603 | -20.118
11.8911 | -34.5608 | -20.1273
11.8972 | -34.5612 | -20.1362
11.9033 | -34.5616 | -20.1449
11.9094 | -34.5621 | -20.1538
11.9155 | -34.5624 | -20.1628
11.9216 | -34.5629 | -20.172
11.9277 | -34.5633 | -20.1809
11.9338 | -34.5637 | -20.1902
11.9399 | -34.5642 | -20.1999
11.946 | -34.5646 | -20.2091
11.9521 | -34.565 | -20.2181
11.9582 | -34.5654 | -20.2272
11.9644 | -34.5659 | -20.2363
11.9705 | -34.5663 | -20.2456
11.9766 | -34.5668 | -20.2548
11.9827 | -34.5672 | -20.2639
11.9888 | -34.5676 | -20.2728
11.9949 | -34.568 | -20.2818
12.001 | -34.5685 | -20.2911
12.0071 | -34.5689 | -20.3
12.0132 | -34.5694 | -20.309
12.0193 | -34.5698 | -20.3186
12.0254 | -34.5702 | -20.3276
12.0315 | -34.5707 | -20.3367
12.0376 | -34.5711 | -20.3455
12.0437 | -34.5716 | -20.3544
12.0498 | -34.572 | -20.3639
12.0559 | -34.5725 | -20.3731
12.062 | -34.5729 | -20.3822
12.0681 | -34.5734 | -20.3913
12.0742 | -34.5738 | -20.4007
12.0803 | -34.5743 | -20.4097
12.0864 | -34.5747 | -20.4188
12.0925 | -34.5752 | -20.4281
12.0986 | -34.5756 | -20.4369
12.1047 | -34.5761 | -20.4458
12.1109 | -34.5766 | -20.4552
12.117 | -34.577 | -20.4642
12.1231 | -34.5775 | -20.4732
12.1292 | -34.5779 | -20.4825
12.1353 | -34.5784 | -20.4914
12.1414 | -34.5789 | -20.5007
12.1475 | -34.5793 | -20.5099
12.1536 | -34.5797 | -20.519
12.1597 | -34.5801 | -20.5281
12.1658 | -34.5806 | -20.5372
12.1719 | -34.581 | -20.5464
12.178 | -34.5815 | -20.5555
12.1841 | -34.5819 | -20.5647
12.1902 | -34.5824 | -20.5738
12.1963 | -34.5828 | -20.5828
12.2024 | -34.5833 | -20.5917
12.2085 | -34.5837 | -20.6009
12.2146 | -34.5842 | -20.6099
12.2207 | -34.5846 | -20.6189
12.2268 | -34.585 | -20.628
12.2329 | -34.5855 | -20.6371
12.239 | -34.5859 | -20.6458
12.2451 | -34.5864 | -20.6547
12.2513 | -34.5868 | -20.6639
12.2574 | -34.5873 | -20.6733
12.2635 | -34.5877 | -20.6819
12.2696 | -34.5881 | -20.6907
12.2757 | -34.5886 | -20.6995
12.2818 | -34.589 | -20.7085
12.2879 | -34.5894 | -20.7172
12.294 | -34.5899 | -20.7262
12.3001 | -34.5903 | -20.7355
12.3062 | -34.5908 | -20.7448
12.3123 | -34.5913 | -20.7538
12.3184 | -34.5917 | -20.7626
12.3245 | -34.5922 | -20.7718
12.3306 | -34.5927 | -20.7805
12.3367 | -34.5931 | -20.7893
12.3428 | -34.5936 | -20.7982
12.3489 | -34.594 | -20.8069
12.355 | -34.5945 | -20.816
12.3611 | -34.5949 | -20.8249
12.3672 | -34.5954 | -20.8338
12.3733 | -34.5958 | -20.8431
12.3794 | -34.5962 | -20.8515
12.3855 | -34.5967 | -20.8604
12.3916 | -34.5971 | -20.8696
12.3978 | -34.5976 | -20.879
12.4039 | -34.598 | -20.8879
12.41 | -34.5985 | -20.8971
12.4161 | -34.599 | -20.9058
12.4222 | -34.5994 | -20.9148
12.4283 | -34.5999 | -20.9239
12.4344 | -34.6003 | -20.9329
12.4405 | -34.6008 | -20.9418
12.4466 | -34.6012 | -20.9506
12.4527 | -34.6016 | -20.9598
12.4588 | -34.6021 | -20.9686
12.4649 | -34.6025 | -20.9778
12.471 | -34.603 | -20.9873
12.4771 | -34.6034 | -20.9956
12.4832 | -34.6039 | -21.0045
12.4893 | -34.6044 | -21.0132
12.4954 | -34.6049 | -21.0221
12.5015 | -34.6054 | -21.0308
12.5076 | -34.6058 | -21.0399
12.5137 | -34.6062 | -21.0492
12.5198 | -34.6067 | -21.0576
12.5259 | -34.6072 | -21.0664
12.532 | -34.6076 | -21.0755
12.5382 | -34.6081 | -21.0842
12.5443 | -34.6085 | -21.0931
12.5504 | -34.6089 | -21.1021
12.5565 | -34.6095 | -21.1108
12.5626 | -34.6099 | -21.1199
12.5687 | -34.6104 | -21.129
12.5748 | -34.6108 | -21.1378
12.5809 | -34.6112 | -21.1463
12.587 | -34.6117 | -21.1556
12.5931 | -34.6122 | -21.1646
12.5992 | -34.6126 | -21.1737
12.6053 | -34.613 | -21.1825
12.6114 | -34.6135 | -21.1913
12.6175 | -34.614 | -21.2003
12.6236 | -34.6144 | -21.209
12.6297 | -34.6149 | -21.2179
12.6358 | -34.6154 | -21.2265
12.6419 | -34.6159 | -21.2357
12.648 | -34.6163 | -21.2449
12.6541 | -34.6168 | -21.2535
12.6602 | -34.6173 | -21.2617
12.6663 | -34.6176 | -21.2704
12.6724 | -34.6181 | -21.2793
12.6785 | -34.6186 | -21.2885
12.6847 | -34.619 | -21.2973
12.6908 | -34.6195 | -21.3061
12.6969 | -34.62 | -21.3148
12.703 | -34.6204 | -21.3237
12.7091 | -34.6208 | -21.3326
12.7152 | -34.6213 | -21.3413
12.7213 | -34.6217 | -21.3502
12.7274 | -34.6222 | -21.3589
12.7335 | -34.6226 | -21.3671
12.7396 | -34.6231 | -21.3755
12.7457 | -34.6235 | -21.3843
12.7518 | -34.6239 | -21.393
12.7579 | -34.6244 | -21.4021
12.764 | -34.6248 | -21.411
12.7701 | -34.6252 | -21.4192
12.7762 | -34.6256 | -21.428
12.7823 | -34.6261 | -21.4369
12.7884 | -34.6265 | -21.4456
12.7945 | -34.6269 | -21.454
12.8006 | -34.6274 | -21.4622
12.8067 | -34.6279 | -21.4713
12.8128 | -34.6283 | -21.4796
12.8189 | -34.6287 | -21.4881
12.8251 | -34.6291 | -21.4967
12.8312 | -34.6295 | -21.5055
12.8373 | -34.63 | -21.5141
12.8434 | -34.6304 | -21.5227
12.8495 | -34.6308 | -21.5312
12.8556 | -34.6313 | -21.54
12.8617 | -34.6317 | -21.5483
12.8678 | -34.6321 | -21.557
12.8739 | -34.6326 | -21.5653
12.88 | -34.633 | -21.5741
12.8861 | -34.6334 | -21.583
12.8922 | -34.6339 | -21.5917
12.8983 | -34.6344 | -21.6001
12.9044 | -34.6348 | -21.6087
12.9105 | -34.6353 | -21.6172
12.9166 | -34.6357 | -21.6262
12.9227 | -34.6361 | -21.6353
12.9288 | -34.6366 | -21.6442
12.9349 | -34.6371 | -21.6531
12.941 | -34.6376 | -21.662
12.9471 | -34.6381 | -21.6706
12.9532 | -34.6386 | -21.679
12.9593 | -34.639 | -21.6879
12.9654 | -34.6394 | -21.6967
12.9716 | -34.6398 | -21.7054
12.9777 | -34.6403 | -21.7135
12.9838 | -34.6407 | -21.7219
12.9899 | -34.6411 | -21.7307
12.996 | -34.6415 | -21.7395
13.0021 | -34.6419 | -21.7482
13.0082 | -34.6423 | -21.7567
13.0143 | -34.6427 | -21.7647
13.0204 | -34.6431 | -21.7731
13.0265 | -34.6436 | -21.7813
13.0326 | -34.644 | -21.7902
13.0387 | -34.6445 | -21.7988
13.0448 | -34.6449 | -21.8075
13.0509 | -34.6454 | -21.8159
13.057 | -34.6458 | -21.8241
13.0631 | -34.6464 | -21.8324
13.0692 | -34.6468 | -21.8408
13.0753 | -34.6473 | -21.8492
13.0814 | -34.6477 | -21.8576
13.0875 | -34.6482 | -21.8657
13.0936 | -34.6486 | -21.8742
13.0997 | -34.6491 | -21.8824
13.1058 | -34.6495 | -21.8904
13.112 | -34.6499 | -21.899
13.1181 | -34.6504 | -21.9079
13.1242 | -34.6508 | -21.9164
13.1303 | -34.6513 | -21.9251
13.1364 | -34.6516 | -21.9338
13.1425 | -34.6521 | -21.9421
13.1486 | -34.6525 | -21.9505
13.1547 | -34.6529 | -21.9588
13.1608 | -34.6534 | -21.9672
13.1669 | -34.6538 | -21.9755
13.173 | -34.6543 | -21.9841
13.1791 | -34.6547 | -21.9925
13.1852 | -34.6551 | -22.0014
13.1913 | -34.6555 | -22.0098
13.1974 | -34.6559 | -22.0183
13.2035 | -34.6563 | -22.0268
13.2096 | -34.6568 | -22.0353
13.2157 | -34.6573 | -22.0436
13.2218 | -34.6577 | -22.052
13.2279 | -34.6582 | -22.0603
13.234 | -34.6586 | -22.0686
13.2401 | -34.659 | -22.0771
13.2462 | -34.6594 | -22.0854
13.2524 | -34.6599 | -22.0938
13.2585 | -34.6603 | -22.1022
13.2646 | -34.6607 | -22.1109
13.2707 | -34.6612 | -22.1193
13.2768 | -34.6616 | -22.1282
13.2829 | -34.662 | -22.1368
13.289 | -34.6624 | -22.1451
13.2951 | -34.6629 | -22.1532
13.3012 | -34.6633 | -22.1617
13.3073 | -34.6637 | -22.1705
13.3134 | -34.6641 | -22.1787
13.3195 | -34.6646 | -22.1868
13.3256 | -34.665 | -22.1953
13.3317 | -34.6654 | -22.2033
13.3378 | -34.6658 | -22.2119
13.3439 | -34.6663 | -22.2204
13.35 | -34.6667 | -22.2287
13.3561 | -34.6671 | -22.2371
13.3622 | -34.6675 | -22.2455
13.3683 | -34.668 | -22.2536
13.3744 | -34.6684 | -22.262
13.3805 | -34.6687 | -22.2703
13.3866 | -34.6692 | -22.2784
13.3927 | -34.6696 | -22.2867
13.3989 | -34.67 | -22.2944
13.405 | -34.6705 | -22.3026
13.4111 | -34.671 | -22.311
13.4172 | -34.6713 | -22.3189
13.4233 | -34.6717 | -22.3273
13.4294 | -34.6721 | -22.3351
13.4355 | -34.6726 | -22.3441
13.4416 | -34.673 | -22.3524
13.4477 | -34.6734 | -22.3608
13.4538 | -34.6738 | -22.3692
13.4599 | -34.6743 | -22.3777
13.466 | -34.6747 | -22.3863
13.4721 | -34.6752 | -22.3946
13.4782 | -34.6756 | -22.4027
13.4843 | -34.6761 | -22.411
13.4904 | -34.6764 | -22.4188
13.4965 | -34.6768 | -22.4275
13.5026 | -34.6772 | -22.4358
13.5087 | -34.6776 | -22.4437
13.5148 | -34.678 | -22.4523
13.5209 | -34.6784 | -22.4604
13.527 | -34.6788 | -22.4691
13.5331 | -34.6792 | -22.4774
13.5393 | -34.6797 | -22.4857
13.5454 | -34.68 | -22.4938
13.5515 | -34.6805 | -22.5022
13.5576 | -34.6809 | -22.5106
13.5637 | -34.6814 | -22.5186
13.5698 | -34.6818 | -22.527
13.5759 | -34.6821 | -22.5349
13.582 | -34.6825 | -22.5432
13.5881 | -34.6829 | -22.5516
13.5942 | -34.6833 | -22.56
13.6003 | -34.6838 | -22.5684
13.6064 | -34.6841 | -22.5768
13.6125 | -34.6846 | -22.5855
13.6186 | -34.685 | -22.5936
13.6247 | -34.6854 | -22.6019
13.6308 | -34.6858 | -22.61
13.6369 | -34.6862 | -22.6182
13.643 | -34.6867 | -22.6265
13.6491 | -34.6871 | -22.6346
13.6552 | -34.6875 | -22.6427
13.6613 | -34.6879 | -22.6513
13.6674 | -34.6883 | -22.6598
13.6735 | -34.6887 | -22.6683
13.6796 | -34.6892 | -22.6763
13.6858 | -34.6895 | -22.6847
13.6919 | -34.69 | -22.693
13.698 | -34.6904 | -22.7013
13.7041 | -34.6908 | -22.7101
13.7102 | -34.6912 | -22.7184
13.7163 | -34.6917 | -22.7263
13.7224 | -34.6921 | -22.7345
13.7285 | -34.6925 | -22.7429
13.7346 | -34.6929 | -22.7511
13.7407 | -34.6933 | -22.7593
13.7468 | -34.6937 | -22.7678
13.7529 | -34.6941 | -22.7758
13.759 | -34.6945 | -22.7841
13.7651 | -34.6949 | -22.7926
13.7712 | -34.6953 | -22.801
13.7773 | -34.6957 | -22.8094
13.7834 | -34.6961 | -22.8177
13.7895 | -34.6966 | -22.8258
13.7956 | -34.697 | -22.8342
13.8017 | -34.6974 | -22.842
13.8078 | -34.6979 | -22.8501
13.8139 | -34.6983 | -22.8582
13.82 | -34.6987 | -22.8662
13.8262 | -34.6991 | -22.8749
13.8323 | -34.6995 | -22.8835
13.8384 | -34.7 | -22.8921
13.8445 | -34.7004 | -22.9007
13.8506 | -34.7008 | -22.9091
13.8567 | -34.7012 | -22.9176
13.8628 | -34.7016 | -22.926
13.8689 | -34.7019 | -22.9344
13.875 | -34.7024 | -22.9426
13.8811 | -34.7028 | -22.9509
13.8872 | -34.7033 | -22.9592
13.8933 | -34.7037 | -22.9673
13.8994 | -34.7041 | -22.9754
13.9055 | -34.7045 | -22.9835
13.9116 | -34.705 | -22.9916
13.9177 | -34.7054 | -22.9994
13.9238 | -34.7059 | -23.0079
13.9299 | -34.7063 | -23.0163
13.936 | -34.7066 | -23.0243
13.9421 | -34.707 | -23.0321
13.9482 | -34.7074 | -23.0405
13.9543 | -34.7078 | -23.0485
13.9604 | -34.7082 | -23.0569
13.9665 | -34.7087 | -23.0653
13.9727 | -34.709 | -23.0734
13.9788 | -34.7094 | -23.0819
13.9849 | -34.7097 | -23.0904
13.991 | -34.7101 | -23.099
13.9971 | -34.7105 | -23.1075
14.0032 | -34.7109 | -23.116
14.0093 | -34.7113 | -23.1241
14.0154 | -34.7116 | -23.1326
14.0215 | -34.712 | -23.1407
14.0276 | -34.7124 | -23.1485
14.0337 | -34.7129 | -23.157
14.0398 | -34.7133 | -23.1651
14.0459 | -34.7136 | -23.1736
14.052 | -34.714 | -23.1819
14.0581 | -34.7144 | -23.1899
14.0642 | -34.7148 | -23.1984
14.0703 | -34.7152 | -23.2067
14.0764 | -34.7156 | -23.2153
14.0825 | -34.716 | -23.2236
14.0886 | -34.7165 | -23.2316
14.0947 | -34.7168 | -23.2397
14.1008 | -34.7173 | -23.2481
14.1069 | -34.7177 | -23.2563
14.1131 | -34.7181 | -23.2649
14.1192 | -34.7184 | -23.2726
14.1253 | -34.7188 | -23.2807
14.1314 | -34.7193 | -23.2888
14.1375 | -34.7197 | -23.2973
14.1436 | -34.7202 | -23.3055
14.1497 | -34.7206 | -23.3135
14.1558 | -34.721 | -23.3219
14.1619 | -34.7214 | -23.33
14.168 | -34.7218 | -23.338
14.1741 | -34.7222 | -23.3461
14.1802 | -34.7226 | -23.3542
14.1863 | -34.723 | -23.3624
14.1924 | -34.7234 | -23.3708
14.1985 | -34.7238 | -23.3788
14.2046 | -34.7242 | -23.387
14.2107 | -34.7246 | -23.3948
14.2168 | -34.7251 | -23.4034
14.2229 | -34.7255 | -23.4116
14.229 | -34.7259 | -23.4199
14.2351 | -34.7263 | -23.4278
14.2412 | -34.7267 | -23.4359
14.2473 | -34.7272 | -23.4442
14.2534 | -34.7276 | -23.4524
14.2596 | -34.728 | -23.4607
14.2657 | -34.7284 | -23.4691
14.2718 | -34.7288 | -23.4774
14.2779 | -34.7292 | -23.4855
14.284 | -34.7295 | -23.4937
14.2901 | -34.73 | -23.5018
14.2962 | -34.7305 | -23.5097
14.3023 | -34.7308 | -23.518
14.3084 | -34.7313 | -23.5266
14.3145 | -34.7317 | -23.5348
14.3206 | -34.7321 | -23.543
14.3267 | -34.7325 | -23.5512
14.3328 | -34.7329 | -23.5593
14.3389 | -34.7333 | -23.5672
14.345 | -34.7337 | -23.5756
14.3511 | -34.7341 | -23.5838
14.3572 | -34.7344 | -23.5922
14.3633 | -34.7348 | -23.6003
14.3694 | -34.7352 | -23.6086
14.3755 | -34.7356 | -23.6165
14.3816 | -34.7361 | -23.6247
14.3877 | -34.7365 | -23.6328
14.3938 | -34.7369 | -23.6413
14.4 | -34.7373 | -23.6494
14.4061 | -34.7377 | -23.6576
14.4122 | -34.7381 | -23.6659
14.4183 | -34.7384 | -23.674
14.4244 | -34.7388 | -23.6821
14.4305 | -34.7392 | -23.6906
14.4366 | -34.7396 | -23.699
14.4427 | -34.7401 | -23.7075
14.4488 | -34.7404 | -23.7154
14.4549 | -34.7409 | -23.7237
14.461 | -34.7413 | -23.732
14.4671 | -34.7417 | -23.7403
14.4732 | -34.7421 | -23.7481
14.4793 | -34.7425 | -23.7569
14.4854 | -34.743 | -23.7647
14.4915 | -34.7434 | -23.7727
14.4976 | -34.7438 | -23.7811
14.5037 | -34.7442 | -23.7894
14.5098 | -34.7446 | -23.7974
14.5159 | -34.745 | -23.8059
14.522 | -34.7454 | -23.8142
14.5281 | -34.7459 | -23.8222
14.5342 | -34.7462 | -23.8305
14.5403 | -34.7467 | -23.8388
14.5465 | -34.7471 | -23.8475
14.5526 | -34.7475 | -23.856
14.5587 | -34.7479 | -23.8642
14.5648 | -34.7483 | -23.8722
14.5709 | -34.7487 | -23.8806
14.577 | -34.7491 | -23.889
14.5831 | -34.7496 | -23.8974
14.5892 | -34.7501 | -23.9052
14.5953 | -34.7504 | -23.9131
14.6014 | -34.7509 | -23.9213
14.6075 | -34.7513 | -23.9291
14.6136 | -34.7517 | -23.9372
14.6197 | -34.7521 | -23.9453
14.6258 | -34.7525 | -23.9536
14.6319 | -34.7529 | -23.9616
14.638 | -34.7534 | -23.9698
14.6441 | -34.7538 | -23.9777
14.6502 | -34.7541 | -23.9856
14.6563 | -34.7545 | -23.9936
14.6624 | -34.755 | -24.0018
14.6685 | -34.7554 | -24.0097
14.6746 | -34.7558 | -24.0181
14.6807 | -34.7562 | -24.0263
14.6869 | -34.7566 | -24.0344
14.693 | -34.757 | -24.0425
14.6991 | -34.7574 | -24.0504
14.7052 | -34.7579 | -24.0586
14.7113 | -34.7583 | -24.0668
14.7174 | -34.7587 | -24.0749
14.7235 | -34.7591 | -24.0828
14.7296 | -34.7595 | -24.0909
14.7357 | -34.76 | -24.099
14.7418 | -34.7604 | -24.1068
14.7479 | -34.7608 | -24.1151
14.754 | -34.7612 | -24.123
14.7601 | -34.7617 | -24.1314
14.7662 | -34.7621 | -24.1395
14.7723 | -34.7625 | -24.1477
14.7784 | -34.7629 | -24.156
14.7845 | -34.7633 | -24.1642
14.7906 | -34.7637 | -24.1725
14.7967 | -34.7641 | -24.1806
14.8028 | -34.7645 | -24.1886
14.8089 | -34.765 | -24.1969
14.815 | -34.7654 | -24.2049
14.8211 | -34.7657 | -24.213
14.8272 | -34.7661 | -24.2214
14.8334 | -34.7666 | -24.2294
14.8395 | -34.767 | -24.2376
14.8456 | -34.7675 | -24.2457
14.8517 | -34.7679 | -24.254
14.8578 | -34.7683 | -24.2619
14.8639 | -34.7687 | -24.2702
14.87 | -34.7692 | -24.2784
14.8761 | -34.7696 | -24.2863
14.8822 | -34.77 | -24.2944
14.8883 | -34.7704 | -24.3023
14.8944 | -34.7709 | -24.3105
14.9005 | -34.7713 | -24.3184
14.9066 | -34.7717 | -24.3266
14.9127 | -34.7722 | -24.3345
14.9188 | -34.7726 | -24.3428
14.9249 | -34.773 | -24.351
14.931 | -34.7734 | -24.3588
14.9371 | -34.7738 | -24.367
14.9432 | -34.7743 | -24.3753
14.9493 | -34.7747 | -24.3835
14.9554 | -34.7751 | -24.3917
14.9615 | -34.7755 | -24.3999
14.9676 | -34.7759 | -24.408
14.9738 | -34.7763 | -24.4162
14.9799 | -34.7768 | -24.4247
14.986 | -34.7772 | -24.4326
14.9921 | -34.7777 | -24.4404
14.9982 | -34.7781 | -24.4488
15.0043 | -34.7785 | -24.4569
15.0104 | -34.7789 | -24.465
15.0165 | -34.7794 | -24.4732
15.0226 | -34.7798 | -24.481
15.0287 | -34.7802 | -24.4896
15.0348 | -34.7807 | -24.4979
15.0409 | -34.7811 | -24.5062
15.047 | -34.7815 | -24.5147
15.0531 | -34.7821 | -24.523
15.0592 | -34.7825 | -24.5313
15.0653 | -34.783 | -24.5393
15.0714 | -34.7833 | -24.5472
15.0775 | -34.7838 | -24.5549
15.0836 | -34.7842 | -24.5631
15.0897 | -34.7847 | -24.5713
15.0958 | -34.7851 | -24.5792
15.1019 | -34.7855 | -24.5873
15.108 | -34.7859 | -24.5956
15.1141 | -34.7863 | -24.6037
15.1203 | -34.7868 | -24.6118
15.1264 | -34.7873 | -24.6202
15.1325 | -34.7878 | -24.6282
15.1386 | -34.7882 | -24.6364
15.1447 | -34.7887 | -24.6445
15.1508 | -34.7892 | -24.6529
15.1569 | -34.7896 | -24.6613
15.163 | -34.7901 | -24.669
15.1691 | -34.7906 | -24.6772
15.1752 | -34.791 | -24.6853
15.1813 | -34.7914 | -24.6934
15.1874 | -34.7919 | -24.7015
15.1935 | -34.7924 | -24.7094
15.1996 | -34.7929 | -24.7177
15.2057 | -34.7933 | -24.7256
15.2118 | -34.7937 | -24.7338
15.2179 | -34.7942 | -24.7415
15.224 | -34.7946 | -24.7497
15.2301 | -34.7951 | -24.7577
15.2362 | -34.7956 | -24.7658
15.2423 | -34.796 | -24.7741
15.2484 | -34.7964 | -24.7822
15.2545 | -34.7969 | -24.7903
15.2607 | -34.7974 | -24.7983
15.2668 | -34.7979 | -24.8063
15.2729 | -34.7984 | -24.8143
15.279 | -34.7988 | -24.8227
15.2851 | -34.7993 | -24.8309
15.2912 | -34.7997 | -24.8388
15.2973 | -34.8001 | -24.8468
15.3034 | -34.8006 | -24.8548
15.3095 | -34.8011 | -24.8627
15.3156 | -34.8015 | -24.8709
15.3217 | -34.8019 | -24.8786
15.3278 | -34.8024 | -24.8867
15.3339 | -34.8029 | -24.8944
15.34 | -34.8034 | -24.9023
15.3461 | -34.8038 | -24.9104
15.3522 | -34.8043 | -24.9184
15.3583 | -34.8047 | -24.9266
15.3644 | -34.8052 | -24.9347
15.3705 | -34.8057 | -24.9426
15.3766 | -34.8061 | -24.9509
15.3827 | -34.8066 | -24.9587
15.3888 | -34.807 | -24.9666
15.3949 | -34.8075 | -24.9744
15.401 | -34.8079 | -24.9826
15.4072 | -34.8085 | -24.9904
15.4133 | -34.8089 | -24.9984
15.4194 | -34.8094 | -25.0066
15.4255 | -34.8099 | -25.0147
15.4316 | -34.8104 | -25.0228
15.4377 | -34.8108 | -25.0308
15.4438 | -34.8113 | -25.0384
15.4499 | -34.8118 | -25.0463
15.456 | -34.8122 | -25.0543
15.4621 | -34.8127 | -25.0624
15.4682 | -34.8132 | -25.0706
15.4743 | -34.8137 | -25.0784
15.4804 | -34.8141 | -25.0865
15.4865 | -34.8146 | -25.0944
15.4926 | -34.815 | -25.1025
15.4987 | -34.8155 | -25.1108
15.5048 | -34.816 | -25.1188
15.5109 | -34.8165 | -25.1265
15.517 | -34.817 | -25.1346
15.5231 | -34.8175 | -25.1425
15.5292 | -34.8179 | -25.1503
15.5353 | -34.8183 | -25.1584
15.5414 | -34.8188 | -25.1666
15.5476 | -34.8193 | -25.1748
15.5537 | -34.8197 | -25.1828
15.5598 | -34.8202 | -25.1905
15.5659 | -34.8207 | -25.1987
15.572 | -34.8212 | -25.2065
15.5781 | -34.8216 | -25.2146
15.5842 | -34.8222 | -25.2223
15.5903 | -34.8226 | -25.2301
15.5964 | -34.8231 | -25.2377
15.6025 | -34.8236 | -25.2455
15.6086 | -34.8241 | -25.2535
15.6147 | -34.8245 | -25.2615
15.6208 | -34.8251 | -25.2692
15.6269 | -34.8256 | -25.277
15.633 | -34.826 | -25.285
15.6391 | -34.8265 | -25.2931
15.6452 | -34.827 | -25.301
15.6513 | -34.8275 | -25.3092
15.6574 | -34.828 | -25.3169
15.6635 | -34.8284 | -25.3249
15.6696 | -34.8289 | -25.3326
15.6757 | -34.8294 | -25.3408
15.6818 | -34.8299 | -25.3489
15.688 | -34.8305 | -25.3569
15.6941 | -34.8309 | -25.3648
15.7002 | -34.8314 | -25.3727
15.7063 | -34.8319 | -25.3807
15.7124 | -34.8323 | -25.3888
15.7185 | -34.8328 | -25.3966
15.7246 | -34.8334 | -25.4044
15.7307 | -34.8338 | -25.4126
15.7368 | -34.8344 | -25.4209
15.7429 | -34.8348 | -25.4286
15.749 | -34.8353 | -25.4366
15.7551 | -34.8358 | -25.4445
15.7612 | -34.8363 | -25.4517
15.7673 | -34.8368 | -25.4596
15.7734 | -34.8372 | -25.4673
15.7795 | -34.8377 | -25.4756
15.7856 | -34.8382 | -25.4831
15.7917 | -34.8387 | -25.4907
15.7978 | -34.8392 | -25.4985
15.8039 | -34.8396 | -25.5064
15.81 | -34.8401 | -25.5142
15.8161 | -34.8406 | -25.5225
15.8222 | -34.841 | -25.5304
15.8283 | -34.8416 | -25.5383
15.8345 | -34.8421 | -25.5462
15.8406 | -34.8426 | -25.5542
15.8467 | -34.8431 | -25.5623
15.8528 | -34.8436 | -25.5701
15.8589 | -34.8441 | -25.578
15.865 | -34.8446 | -25.5859
15.8711 | -34.8451 | -25.5942
15.8772 | -34.8456 | -25.6021
15.8833 | -34.8461 | -25.6106
15.8894 | -34.8465 | -25.6183
15.8955 | -34.847 | -25.6261
15.9016 | -34.8476 | -25.634
15.9077 | -34.8481 | -25.642
15.9138 | -34.8486 | -25.6497
15.9199 | -34.8492 | -25.6576
15.926 | -34.8497 | -25.6658
15.9321 | -34.8502 | -25.6735
15.9382 | -34.8507 | -25.681
15.9443 | -34.8513 | -25.6891
15.9504 | -34.8518 | -25.6972
15.9565 | -34.8523 | -25.7049
15.9626 | -34.8528 | -25.7127
15.9687 | -34.8534 | -25.7203
15.9749 | -34.8539 | -25.7283
15.981 | -34.8543 | -25.7361
15.9871 | -34.8549 | -25.7439
15.9932 | -34.8554 | -25.7518
15.9993 | -34.8559 | -25.7596
16.0054 | -34.8564 | -25.7673
16.0115 | -34.857 | -25.775
16.0176 | -34.8574 | -25.783
16.0237 | -34.858 | -25.7913
16.0298 | -34.8585 | -25.7986
16.0359 | -34.859 | -25.8063
16.042 | -34.8595 | -25.8142
16.0481 | -34.86 | -25.8222
16.0542 | -34.8605 | -25.8303
16.0603 | -34.8611 | -25.838
16.0664 | -34.8615 | -25.8458
16.0725 | -34.8621 | -25.8533
16.0786 | -34.8626 | -25.8615
16.0847 | -34.8631 | -25.869
16.0908 | -34.8636 | -25.8769
16.0969 | -34.8641 | -25.8845
16.103 | -34.8646 | -25.8926
16.1091 | -34.8651 | -25.9003
16.1152 | -34.8656 | -25.9082
16.1214 | -34.8661 | -25.916
16.1275 | -34.8666 | -25.9237
16.1336 | -34.8671 | -25.9316
16.1397 | -34.8677 | -25.9395
16.1458 | -34.8682 | -25.9472
16.1519 | -34.8687 | -25.9551
16.158 | -34.8692 | -25.963
16.1641 | -34.8697 | -25.9708
16.1702 | -34.8702 | -25.9785
16.1763 | -34.8707 | -25.9862
16.1824 | -34.8712 | -25.9937
16.1885 | -34.8718 | -26.0014
16.1946 | -34.8723 | -26.0093
16.2007 | -34.8728 | -26.0177
16.2068 | -34.8733 | -26.0252
16.2129 | -34.8738 | -26.0331
16.219 | -34.8743 | -26.0407
16.2251 | -34.8748 | -26.0485
16.2312 | -34.8753 | -26.0565
16.2373 | -34.8758 | -26.0642
16.2434 | -34.8763 | -26.0722
16.2495 | -34.8769 | -26.0798
16.2556 | -34.8774 | -26.0877
16.2618 | -34.878 | -26.0952
16.2679 | -34.8785 | -26.1029
16.274 | -34.879 | -26.111
16.2801 | -34.8796 | -26.1187
16.2862 | -34.8801 | -26.1263
16.2923 | -34.8806 | -26.1342
16.2984 | -34.8811 | -26.142
16.3045 | -34.8817 | -26.1498
16.3106 | -34.8822 | -26.1573
16.3167 | -34.8827 | -26.1652
16.3228 | -34.8832 | -26.173
16.3289 | -34.8837 | -26.1808
16.335 | -34.8842 | -26.1887
16.3411 | -34.8847 | -26.1963
16.3472 | -34.8852 | -26.2044
16.3533 | -34.8857 | -26.2121
16.3594 | -34.8863 | -26.2194
16.3655 | -34.8868 | -26.2274
16.3716 | -34.8873 | -26.2351
16.3777 | -34.8878 | -26.2428
16.3838 | -34.8884 | -26.2506
16.3899 | -34.8889 | -26.2585
16.396 | -34.8894 | -26.2665
16.4021 | -34.8899 | -26.2742
16.4083 | -34.8905 | -26.2819
16.4144 | -34.891 | -26.2895
16.4205 | -34.8915 | -26.2973
16.4266 | -34.8921 | -26.3048
16.4327 | -34.8926 | -26.3124
16.4388 | -34.8931 | -26.3202
16.4449 | -34.8936 | -26.328
16.451 | -34.8941 | -26.3359
16.4571 | -34.8946 | -26.3439
16.4632 | -34.8951 | -26.3512
16.4693 | -34.8956 | -26.359
16.4754 | -34.8961 | -26.3668
16.4815 | -34.8966 | -26.3747
16.4876 | -34.8972 | -26.3825
16.4937 | -34.8977 | -26.3902
16.4998 | -34.8983 | -26.3976
16.5059 | -34.8989 | -26.4051
16.512 | -34.8994 | -26.4132
16.5181 | -34.8999 | -26.421
16.5242 | -34.9004 | -26.4288
16.5303 | -34.901 | -26.4362
16.5364 | -34.9015 | -26.4438
16.5425 | -34.902 | -26.4516
16.5487 | -34.9025 | -26.4593
16.5548 | -34.9029 | -26.4669
16.5609 | -34.9035 | -26.4749
16.567 | -34.904 | -26.4827
16.5731 | -34.9045 | -26.4904
16.5792 | -34.905 | -26.4982
16.5853 | -34.9056 | -26.5056
16.5914 | -34.9061 | -26.5133
16.5975 | -34.9066 | -26.5211
16.6036 | -34.9071 | -26.5289
16.6097 | -34.9077 | -26.5367
16.6158 | -34.9081 | -26.5445
16.6219 | -34.9087 | -26.5521
16.628 | -34.9092 | -26.5601
16.6341 | -34.9097 | -26.5677
16.6402 | -34.9102 | -26.5751
16.6463 | -34.9107 | -26.5827
16.6524 | -34.9112 | -26.5903
16.6585 | -34.9118 | -26.5982
16.6646 | -34.9123 | -26.6059
16.6707 | -34.9129 | -26.6133
16.6768 | -34.9134 | -26.6207
16.6829 | -34.9139 | -26.6283
16.689 | -34.9144 | -26.6362
16.6952 | -34.9149 | -26.6438
16.7013 | -34.9154 | -26.6515
16.7074 | -34.916 | -26.6592
16.7135 | -34.9165 | -26.667
16.7196 | -34.917 | -26.6747
16.7257 | -34.9175 | -26.6822
16.7318 | -34.918 | -26.69
16.7379 | -34.9186 | -26.6979
16.744 | -34.9191 | -26.7054
16.7501 | -34.9196 | -26.7133
16.7562 | -34.9202 | -26.721
16.7623 | -34.9207 | -26.7287
16.7684 | -34.9212 | -26.7367
16.7745 | -34.9218 | -26.7441
16.7806 | -34.9223 | -26.7516
16.7867 | -34.9229 | -26.7592
16.7928 | -34.9233 | -26.7667
16.7989 | -34.9238 | -26.7741
16.805 | -34.9243 | -26.7817
16.8111 | -34.9249 | -26.7893
16.8172 | -34.9254 | -26.7972
16.8233 | -34.9259 | -26.805
16.8294 | -34.9265 | -26.8126
16.8356 | -34.9269 | -26.8205
16.8417 | -34.9275 | -26.828
16.8478 | -34.928 | -26.8356
16.8539 | -34.9285 | -26.8433
16.86 | -34.9291 | -26.8511
16.8661 | -34.9296 | -26.8589
16.8722 | -34.9301 | -26.8666
16.8783 | -34.9306 | -26.874
16.8844 | -34.9312 | -26.8815
16.8905 | -34.9317 | -26.8893
16.8966 | -34.9322 | -26.8972
16.9027 | -34.9327 | -26.9049
16.9088 | -34.9332 | -26.913
16.9149 | -34.9337 | -26.921
16.921 | -34.9342 | -26.9282
16.9271 | -34.9348 | -26.9362
16.9332 | -34.9353 | -26.944
16.9393 | -34.9357 | -26.9522
16.9454 | -34.9362 | -26.9597
16.9515 | -34.9367 | -26.9677
16.9576 | -34.9373 | -26.9752
16.9637 | -34.9379 | -26.9829
16.9698 | -34.9384 | -26.9903
16.9759 | -34.9389 | -26.9977
16.9821 | -34.9395 | -27.0057
16.9882 | -34.94 | -27.0134
16.9943 | -34.9405 | -27.0213
17.0004 | -34.9411 | -27.029
17.0065 | -34.9416 | -27.0367
17.0126 | -34.9421 | -27.0443
17.0187 | -34.9426 | -27.052
17.0248 | -34.9431 | -27.0594
17.0309 | -34.9436 | -27.0671
17.037 | -34.9442 | -27.0749
17.0431 | -34.9447 | -27.0825
17.0492 | -34.9452 | -27.0903
17.0553 | -34.9457 | -27.098
17.0614 | -34.9462 | -27.1055
17.0675 | -34.9467 | -27.1134
17.0736 | -34.9472 | -27.121
17.0797 | -34.9477 | -27.1282
17.0858 | -34.9482 | -27.1359
17.0919 | -34.9488 | -27.1439
17.098 | -34.9492 | -27.1513
17.1041 | -34.9498 | -27.1589
17.1102 | -34.9503 | -27.1668
17.1163 | -34.9508 | -27.1745
17.1225 | -34.9513 | -27.1819
17.1286 | -34.9518 | -27.1893
17.1347 | -34.9523 | -27.197
17.1408 | -34.9529 | -27.2046
17.1469 | -34.9534 | -27.2123
17.153 | -34.954 | -27.2197
17.1591 | -34.9545 | -27.2272
17.1652 | -34.955 | -27.2347
17.1713 | -34.9555 | -27.2423
17.1774 | -34.956 | -27.25
17.1835 | -34.9566 | -27.2577
17.1896 | -34.957 | -27.2653
17.1957 | -34.9575 | -27.2729
17.2018 | -34.958 | -27.2805
17.2079 | -34.9585 | -27.2884
17.214 | -34.959 | -27.2959
17.2201 | -34.9595 | -27.3035
17.2262 | -34.96 | -27.3112
17.2323 | -34.9605 | -27.3189
17.2384 | -34.9611 | -27.3266
17.2445 | -34.9615 | -27.3342
17.2506 | -34.962 | -27.3417
17.2567 | -34.9625 | -27.3495
17.2628 | -34.9631 | -27.3572
17.269 | -34.9636 | -27.365
17.2751 | -34.9641 | -27.3727
17.2812 | -34.9646 | -27.3804
17.2873 | -34.9651 | -27.3882
17.2934 | -34.9656 | -27.396
17.2995 | -34.9661 | -27.4034
17.3056 | -34.9666 | -27.4113
17.3117 | -34.9671 | -27.419
17.3178 | -34.9675 | -27.4267
17.3239 | -34.968 | -27.4343
17.33 | -34.9685 | -27.4421
17.3361 | -34.969 | -27.4496
17.3422 | -34.9695 | -27.4569
17.3483 | -34.97 | -27.4647
17.3544 | -34.9705 | -27.4722
17.3605 | -34.971 | -27.4797
17.3666 | -34.9715 | -27.4874
17.3727 | -34.972 | -27.495
17.3788 | -34.9726 | -27.5024
17.3849 | -34.9731 | -27.5098
17.391 | -34.9736 | -27.5175
17.3971 | -34.9741 | -27.5249
17.4032 | -34.9746 | -27.5326
17.4094 | -34.9751 | -27.5405
17.4155 | -34.9756 | -27.5482
17.4216 | -34.9761 | -27.5558
17.4277 | -34.9766 | -27.5639
17.4338 | -34.9772 | -27.5717
17.4399 | -34.9776 | -27.5795
17.446 | -34.9781 | -27.587
17.4521 | -34.9787 | -27.5946
17.4582 | -34.9792 | -27.6021
17.4643 | -34.9796 | -27.6098
17.4704 | -34.9801 | -27.6172
17.4765 | -34.9806 | -27.6248
17.4826 | -34.9811 | -27.6325
17.4887 | -34.9816 | -27.6402
17.4948 | -34.9821 | -27.6477
17.5009 | -34.9825 | -27.6552
17.507 | -34.9831 | -27.6628
17.5131 | -34.9835 | -27.6705
17.5192 | -34.984 | -27.6782
17.5253 | -34.9845 | -27.6858
17.5314 | -34.9849 | -27.6934
17.5375 | -34.9855 | -27.7011
17.5436 | -34.9859 | -27.7087
17.5497 | -34.9865 | -27.7166
17.5559 | -34.9869 | -27.724
17.562 | -34.9875 | -27.7314
17.5681 | -34.988 | -27.7391
17.5742 | -34.9885 | -27.7467
17.5803 | -34.9889 | -27.7543
17.5864 | -34.9895 | -27.7616
17.5925 | -34.99 | -27.7697
17.5986 | -34.9905 | -27.7772
17.6047 | -34.9911 | -27.7852
17.6108 | -34.9915 | -27.7927
17.6169 | -34.992 | -27.8002
17.623 | -34.9924 | -27.8077
17.6291 | -34.9929 | -27.8152
17.6352 | -34.9935 | -27.8227
17.6413 | -34.994 | -27.8304
17.6474 | -34.9945 | -27.8381
17.6535 | -34.995 | -27.8457
17.6596 | -34.9954 | -27.8532
17.6657 | -34.9959 | -27.8606
17.6718 | -34.9964 | -27.8683
17.6779 | -34.9968 | -27.8758
17.684 | -34.9974 | -27.8836
17.6901 | -34.9979 | -27.8911
17.6963 | -34.9984 | -27.8987
17.7024 | -34.9989 | -27.9062
17.7085 | -34.9994 | -27.9137
17.7146 | -34.9999 | -27.9213
17.7207 | -35.0004 | -27.9288
17.7268 | -35.0009 | -27.9361
17.7329 | -35.0013 | -27.9441
17.739 | -35.0018 | -27.9518
17.7451 | -35.0024 | -27.9592
17.7512 | -35.0028 | -27.9669
17.7573 | -35.0033 | -27.9745
17.7634 | -35.0038 | -27.9817
17.7695 | -35.0042 | -27.9893
17.7756 | -35.0048 | -27.9969
17.7817 | -35.0053 | -28.0043
17.7878 | -35.0058 | -28.0121
17.7939 | -35.0063 | -28.0197
17.8 | -35.0068 | -28.0274
17.8061 | -35.0072 | -28.0352
17.8122 | -35.0078 | -28.0426
17.8183 | -35.0083 | -28.0502
17.8244 | -35.0088 | -28.0578
17.8305 | -35.0093 | -28.0653
17.8366 | -35.0098 | -28.0725
17.8428 | -35.0102 | -28.0803
17.8489 | -35.0107 | -28.0878
17.855 | -35.0113 | -28.095
17.8611 | -35.0117 | -28.1024
17.8672 | -35.0123 | -28.1099
17.8733 | -35.0128 | -28.1174
17.8794 | -35.0133 | -28.1251
17.8855 | -35.0138 | -28.1325
17.8916 | -35.0142 | -28.1398
17.8977 | -35.0147 | -28.1473
17.9038 | -35.0152 | -28.1548
17.9099 | -35.0157 | -28.1624
17.916 | -35.0162 | -28.1699
17.9221 | -35.0167 | -28.1772
17.9282 | -35.0171 | -28.185
17.9343 | -35.0177 | -28.1927
17.9404 | -35.0182 | -28.2002
17.9465 | -35.0186 | -28.2078
17.9526 | -35.0191 | -28.2156
17.9587 | -35.0196 | -28.2233
17.9648 | -35.0201 | -28.2307
17.9709 | -35.0206 | -28.2382
17.977 | -35.0211 | -28.2459
17.9832 | -35.0215 | -28.2538
17.9893 | -35.022 | -28.2614
17.9954 | -35.0225 | -28.2693
18.0015 | -35.0231 | -28.277
18.0076 | -35.0236 | -28.2839
18.0137 | -35.024 | -28.2917
18.0198 | -35.0245 | -28.299
18.0259 | -35.025 | -28.3068
18.032 | -35.0254 | -28.3142
18.0381 | -35.0259 | -28.3215
18.0442 | -35.0264 | -28.329
18.0503 | -35.0269 | -28.3368
18.0564 | -35.0274 | -28.3443
18.0625 | -35.0279 | -28.3519
18.0686 | -35.0284 | -28.3593
18.0747 | -35.0288 | -28.3669
18.0808 | -35.0293 | -28.3744
18.0869 | -35.0299 | -28.382
18.093 | -35.0304 | -28.3895
18.0991 | -35.0309 | -28.3972
18.1052 | -35.0314 | -28.4051
18.1113 | -35.0319 | -28.4123
18.1174 | -35.0324 | -28.4197
18.1236 | -35.0329 | -28.4273
18.1297 | -35.0334 | -28.4348
18.1358 | -35.034 | -28.4425
18.1419 | -35.0345 | -28.4504
18.148 | -35.035 | -28.4579
18.1541 | -35.0355 | -28.4654
18.1602 | -35.0361 | -28.4731
18.1663 | -35.0366 | -28.4806
18.1724 | -35.0371 | -28.4883
18.1785 | -35.0376 | -28.496
18.1846 | -35.038 | -28.5034
18.1907 | -35.0385 | -28.5112
18.1968 | -35.0389 | -28.5188
18.2029 | -35.0394 | -28.5258
18.209 | -35.0399 | -28.5334
18.2151 | -35.0405 | -28.5411
18.2212 | -35.041 | -28.5486
18.2273 | -35.0415 | -28.5562
18.2334 | -35.042 | -28.5636
18.2395 | -35.0424 | -28.5706
18.2456 | -35.0429 | -28.5782
18.2517 | -35.0434 | -28.5857
18.2578 | -35.0438 | -28.593
18.2639 | -35.0443 | -28.6011
18.2701 | -35.0449 | -28.6085
18.2762 | -35.0453 | -28.6165
18.2823 | -35.0459 | -28.6239
18.2884 | -35.0463 | -28.6315
18.2945 | -35.0469 | -28.639
18.3006 | -35.0474 | -28.6465
18.3067 | -35.0478 | -28.6537
18.3128 | -35.0484 | -28.6613
18.3189 | -35.0488 | -28.6689
18.325 | -35.0494 | -28.6759
18.3311 | -35.0498 | -28.6836
18.3372 | -35.0503 | -28.6909
18.3433 | -35.0508 | -28.6981
18.3494 | -35.0513 | -28.7057
18.3555 | -35.0518 | -28.7131
18.3616 | -35.0523 | -28.7203
18.3677 | -35.0528 | -28.7278
18.3738 | -35.0532 | -28.7353
18.3799 | -35.0537 | -28.7426
18.386 | -35.0541 | -28.7499
18.3921 | -35.0546 | -28.7572
18.3982 | -35.0551 | -28.7645
18.4043 | -35.0555 | -28.7718
18.4105 | -35.0559 | -28.7793
18.4166 | -35.0565 | -28.7867
18.4227 | -35.057 | -28.7945
18.4288 | -35.0575 | -28.8022
18.4349 | -35.058 | -28.8093
18.441 | -35.0585 | -28.8164
18.4471 | -35.059 | -28.8241
18.4532 | -35.0594 | -28.8317
18.4593 | -35.0599 | -28.8392
18.4654 | -35.0604 | -28.8463
18.4715 | -35.0609 | -28.8538
18.4776 | -35.0614 | -28.8614
18.4837 | -35.0619 | -28.8689
18.4898 | -35.0624 | -28.8766
18.4959 | -35.0629 | -28.8841
18.502 | -35.0634 | -28.8919
18.5081 | -35.0639 | -28.8992
18.5142 | -35.0643 | -28.9069
18.5203 | -35.0648 | -28.9145
18.5264 | -35.0653 | -28.9216
18.5325 | -35.0658 | -28.9291
18.5386 | -35.0663 | -28.9374
18.5447 | -35.0668 | -28.945
18.5508 | -35.0673 | -28.9527
18.557 | -35.0678 | -28.9604
18.5631 | -35.0682 | -28.9678
18.5692 | -35.0687 | -28.9747
18.5753 | -35.0692 | -28.9819
18.5814 | -35.0697 | -28.9896
18.5875 | -35.0702 | -28.9973
18.5936 | -35.0707 | -29.0051
18.5997 | -35.0711 | -29.0127
18.6058 | -35.0716 | -29.02
18.6119 | -35.0721 | -29.0273
18.618 | -35.0726 | -29.0346
18.6241 | -35.0731 | -29.0423
18.6302 | -35.0736 | -29.0498
18.6363 | -35.0741 | -29.057
18.6424 | -35.0746 | -29.0641
18.6485 | -35.075 | -29.0714
18.6546 | -35.0755 | -29.0791
18.6607 | -35.076 | -29.0867
18.6668 | -35.0765 | -29.0941
18.6729 | -35.077 | -29.1016
18.679 | -35.0775 | -29.1086
18.6851 | -35.078 | -29.1159
18.6912 | -35.0785 | -29.1234
18.6974 | -35.079 | -29.1308
18.7035 | -35.0794 | -29.1381
18.7096 | -35.0799 | -29.1458
18.7157 | -35.0805 | -29.1532
18.7218 | -35.081 | -29.1605
18.7279 | -35.0815 | -29.1678
18.734 | -35.082 | -29.1752
18.7401 | -35.0825 | -29.1825
18.7462 | -35.083 | -29.1896
18.7523 | -35.0835 | -29.1964
18.7584 | -35.084 | -29.2041
18.7645 | -35.0845 | -29.2111
18.7706 | -35.085 | -29.2187
18.7767 | -35.0855 | -29.2263
18.7828 | -35.086 | -29.2335
18.7889 | -35.0866 | -29.241
18.795 | -35.087 | -29.2485
18.8011 | -35.0876 | -29.256
18.8072 | -35.088 | -29.2634
18.8133 | -35.0885 | -29.2707
18.8194 | -35.0889 | -29.2782
18.8255 | -35.0894 | -29.2861
18.8316 | -35.0899 | -29.294
18.8377 | -35.0903 | -29.3013
18.8439 | -35.0908 | -29.3082
18.85 | -35.0913 | -29.3158
18.8561 | -35.0918 | -29.3229
18.8622 | -35.0923 | -29.3306
18.8683 | -35.0928 | -29.3374
18.8744 | -35.0934 | -29.3447
18.8805 | -35.0938 | -29.3527
18.8866 | -35.0944 | -29.3602
18.8927 | -35.0949 | -29.3673
18.8988 | -35.0954 | -29.3748
18.9049 | -35.0959 | -29.382
18.911 | -35.0964 | -29.3891
18.9171 | -35.0969 | -29.3964
18.9232 | -35.0974 | -29.4038
18.9293 | -35.0979 | -29.4113
18.9354 | -35.0983 | -29.4184
18.9415 | -35.0988 | -29.4256
18.9476 | -35.0992 | -29.433
18.9537 | -35.0997 | -29.4406
18.9598 | -35.1002 | -29.4481
18.9659 | -35.1008 | -29.4556
18.972 | -35.1012 | -29.4632
18.9781 | -35.1017 | -29.4703
18.9843 | -35.1021 | -29.4782
18.9904 | -35.1026 | -29.4854
18.9965 | -35.1031 | -29.4934
19.0026 | -35.1037 | -29.5007
19.0087 | -35.1041 | -29.5077
19.0148 | -35.1047 | -29.5154
19.0209 | -35.1051 | -29.5232
19.027 | -35.1056 | -29.5307
19.0331 | -35.1061 | -29.5382
19.0392 | -35.1066 | -29.5458
19.0453 | -35.1071 | -29.5536
19.0514 | -35.1075 | -29.5609
19.0575 | -35.108 | -29.5678
19.0636 | -35.1085 | -29.5744
19.0697 | -35.1091 | -29.5823
19.0758 | -35.1096 | -29.5901
19.0819 | -35.1101 | -29.5981
19.088 | -35.1105 | -29.6051
19.0941 | -35.111 | -29.6117
19.1002 | -35.1113 | -29.6189
19.1063 | -35.1118 | -29.6263
19.1124 | -35.1123 | -29.6339
19.1185 | -35.1128 | -29.6412
19.1246 | -35.1132 | -29.6482
19.1308 | -35.1138 | -29.6554
19.1369 | -35.1143 | -29.6626
19.143 | -35.1148 | -29.6698
19.1491 | -35.1153 | -29.6769
19.1552 | -35.1158 | -29.6846
19.1613 | -35.1163 | -29.692
19.1674 | -35.1168 | -29.6994
19.1735 | -35.1173 | -29.7062
19.1796 | -35.1178 | -29.713
19.1857 | -35.1183 | -29.72
19.1918 | -35.1188 | -29.7275
19.1979 | -35.1193 | -29.7351
19.204 | -35.1197 | -29.7429
19.2101 | -35.1202 | -29.75
19.2162 | -35.1207 | -29.7572
19.2223 | -35.1212 | -29.7644
19.2284 | -35.1218 | -29.7717
19.2345 | -35.1223 | -29.7789
19.2406 | -35.1228 | -29.786
19.2467 | -35.1232 | -29.7934
19.2528 | -35.1237 | -29.8008
19.2589 | -35.1242 | -29.808
19.265 | -35.1246 | -29.8155
19.2712 | -35.125 | -29.8231
19.2773 | -35.1256 | -29.8302
19.2834 | -35.1261 | -29.8373
19.2895 | -35.1266 | -29.8446
19.2956 | -35.127 | -29.8519
19.3017 | -35.1276 | -29.8591
19.3078 | -35.128 | -29.8661
19.3139 | -35.1285 | -29.8729
19.32 | -35.129 | -29.8804
19.3261 | -35.1294 | -29.8878
19.3322 | -35.1299 | -29.8949
19.3383 | -35.1305 | -29.9021
19.3444 | -35.131 | -29.9094
19.3505 | -35.1314 | -29.9166
19.3566 | -35.1319 | -29.924
19.3627 | -35.1325 | -29.9315
19.3688 | -35.133 | -29.9389
19.3749 | -35.1335 | -29.9468
19.381 | -35.1339 | -29.9537
19.3871 | -35.1344 | -29.9614
19.3932 | -35.1349 | -29.9688
19.3993 | -35.1354 | -29.9764
19.4054 | -35.1359 | -29.9843
19.4115 | -35.1364 | -29.9916
19.4177 | -35.1368 | -29.9986
19.4238 | -35.1373 | -30.0054
19.4299 | -35.1377 | -30.0123
19.436 | -35.1383 | -30.0197
19.4421 | -35.1388 | -30.0263
19.4482 | -35.1393 | -30.0336
19.4543 | -35.1399 | -30.0401
19.4604 | -35.1403 | -30.0472
19.4665 | -35.1407 | -30.0542
19.4726 | -35.1412 | -30.0614
19.4787 | -35.1418 | -30.0685
19.4848 | -35.1422 | -30.076
19.4909 | -35.1428 | -30.083
19.497 | -35.1433 | -30.0903
19.5031 | -35.1437 | -30.0972
19.5092 | -35.1442 | -30.1043
19.5153 | -35.1448 | -30.1114
19.5214 | -35.1453 | -30.1174
19.5275 | -35.1458 | -30.1246
19.5336 | -35.1462 | -30.1312
19.5397 | -35.1467 | -30.1385
19.5458 | -35.1472 | -30.1456
19.5519 | -35.1477 | -30.1529
19.5581 | -35.1483 | -30.1601
19.5642 | -35.1488 | -30.1669
19.5703 | -35.1493 | -30.1741
19.5764 | -35.1497 | -30.1815
19.5825 | -35.1501 | -30.1885
19.5886 | -35.1505 | -30.1958
19.5947 | -35.151 | -30.2023
19.6008 | -35.1514 | -30.209
19.6069 | -35.1519 | -30.2157
19.613 | -35.1524 | -30.2234
19.6191 | -35.1529 | -30.231
19.6252 | -35.1534 | -30.2378
19.6313 | -35.1538 | -30.2445
19.6374 | -35.1542 | -30.252
19.6435 | -35.1548 | -30.2587
19.6496 | -35.1553 | -30.2657
19.6557 | -35.1557 | -30.2724
19.6618 | -35.1561 | -30.2793
19.6679 | -35.1565 | -30.2865
19.674 | -35.157 | -30.2935
19.6801 | -35.1575 | -30.3007
19.6862 | -35.158 | -30.3071
19.6923 | -35.1585 | -30.3142
19.6984 | -35.1589 | -30.3208
19.7046 | -35.1595 | -30.3278
19.7107 | -35.16 | -30.3351
19.7168 | -35.1604 | -30.3425
19.7229 | -35.1609 | -30.3491
19.729 | -35.1613 | -30.356
19.7351 | -35.1618 | -30.3631
19.7412 | -35.1623 | -30.3698
19.7473 | -35.1628 | -30.3775
19.7534 | -35.1633 | -30.3844
19.7595 | -35.1637 | -30.3913
19.7656 | -35.1642 | -30.3987
19.7717 | -35.1647 | -30.4051
19.7778 | -35.1651 | -30.4127
19.7839 | -35.1657 | -30.4193
19.79 | -35.1662 | -30.4268
19.7961 | -35.1667 | -30.4339
19.8022 | -35.1672 | -30.4411
19.8083 | -35.1678 | -30.4481
19.8144 | -35.1683 | -30.4555
19.8205 | -35.1687 | -30.4627
19.8266 | -35.1692 | -30.47
19.8327 | -35.1696 | -30.4772
19.8388 | -35.1702 | -30.4843
19.845 | -35.1707 | -30.4911
19.8511 | -35.1712 | -30.4983
19.8572 | -35.1716 | -30.5054
19.8633 | -35.1721 | -30.5124
19.8694 | -35.1726 | -30.5192
19.8755 | -35.173 | -30.5258
19.8816 | -35.1735 | -30.5329
19.8877 | -35.174 | -30.5404
19.8938 | -35.1744 | -30.5472
19.8999 | -35.1749 | -30.5542
19.906 | -35.1754 | -30.5613
19.9121 | -35.1759 | -30.5677
19.9182 | -35.1764 | -30.5743
19.9243 | -35.1769 | -30.5815
19.9304 | -35.1774 | -30.5888
19.9365 | -35.1779 | -30.5962
19.9426 | -35.1784 | -30.6028
19.9487 | -35.1789 | -30.6095
19.9548 | -35.1794 | -30.6172
19.9609 | -35.1799 | -30.6241
19.967 | -35.1804 | -30.6308
19.9731 | -35.1809 | -30.6376
19.9792 | -35.1814 | -30.6438
19.9853 | -35.182 | -30.651
19.9915 | -35.1824 | -30.6577
19.9976 | -35.1829 | -30.6646
20.0037 | -35.1834 | -30.6722
20.0098 | -35.1838 | -30.6789
20.0159 | -35.1844 | -30.6857
20.022 | -35.185 | -30.6931
20.0281 | -35.1855 | -30.7001
20.0342 | -35.1859 | -30.7068
20.0403 | -35.1863 | -30.713
20.0464 | -35.1868 | -30.72
20.0525 | -35.1874 | -30.7267
20.0586 | -35.1879 | -30.7337
20.0647 | -35.1883 | -30.7414
20.0708 | -35.1888 | -30.7489
20.0769 | -35.1893 | -30.7554
20.083 | -35.1897 | -30.7624
20.0891 | -35.1901 | -30.7692
20.0952 | -35.1906 | -30.7759
20.1013 | -35.191 | -30.783
20.1074 | -35.1915 | -30.7904
20.1135 | -35.1919 | -30.7967
20.1196 | -35.1924 | -30.8034
20.1257 | -35.1929 | -30.8108
20.1319 | -35.1934 | -30.8177
20.138 | -35.1938 | -30.8251
20.1441 | -35.1944 | -30.8319
20.1502 | -35.1948 | -30.839
20.1563 | -35.1953 | -30.8451
20.1624 | -35.1958 | -30.8517
20.1685 | -35.1963 | -30.8591
20.1746 | -35.1968 | -30.8657
20.1807 | -35.1972 | -30.873
20.1868 | -35.1977 | -30.88
20.1929 | -35.1982 | -30.8862
20.199 | -35.1986 | -30.893
20.2051 | -35.1991 | -30.8997
20.2112 | -35.1997 | -30.9061
20.2173 | -35.2001 | -30.9133
20.2234 | -35.2006 | -30.9201
20.2295 | -35.2012 | -30.9276
20.2356 | -35.2017 | -30.9347
20.2417 | -35.2022 | -30.9414
20.2478 | -35.2025 | -30.9481
20.2539 | -35.2031 | -30.9548
20.26 | -35.2035 | -30.9611
20.2661 | -35.204 | -30.9686
20.2723 | -35.2044 | -30.9753
20.2784 | -35.205 | -30.982
20.2845 | -35.2054 | -30.9892
20.2906 | -35.2059 | -30.9963
20.2967 | -35.2063 | -31.0021
20.3028 | -35.2068 | -31.0084
20.3089 | -35.2072 | -31.0153
20.315 | -35.2077 | -31.022
20.3211 | -35.2082 | -31.0292
20.3272 | -35.2088 | -31.0361
20.3333 | -35.2092 | -31.0426
20.3394 | -35.2096 | -31.0489
20.3455 | -35.2101 | -31.056
20.3516 | -35.2106 | -31.063
20.3577 | -35.2111 | -31.0701
20.3638 | -35.2116 | -31.0772
20.3699 | -35.2121 | -31.084
20.376 | -35.2126 | -31.0913
20.3821 | -35.213 | -31.0985
20.3882 | -35.2135 | -31.1061
20.3943 | -35.214 | -31.1128
20.4004 | -35.2145 | -31.1195
20.4065 | -35.2149 | -31.1261
20.4126 | -35.2154 | -31.1331
20.4188 | -35.2158 | -31.1399
20.4249 | -35.2163 | -31.1472
20.431 | -35.2166 | -31.1537
20.4371 | -35.217 | -31.1597
20.4432 | -35.2175 | -31.1669
20.4493 | -35.218 | -31.1732
20.4554 | -35.2185 | -31.1793
20.4615 | -35.2189 | -31.1863
20.4676 | -35.2193 | -31.1938
20.4737 | -35.2198 | -31.2001
20.4798 | -35.2203 | -31.2079
20.4859 | -35.2207 | -31.2142
20.492 | -35.2211 | -31.2211
20.4981 | -35.2216 | -31.2282
20.5042 | -35.2221 | -31.2341
20.5103 | -35.2225 | -31.2403
20.5164 | -35.2231 | -31.2472
20.5225 | -35.2234 | -31.2543
20.5286 | -35.2239 | -31.2613
20.5347 | -35.2244 | -31.2677
20.5408 | -35.2249 | -31.2745
20.5469 | -35.2253 | -31.2819
20.553 | -35.2256 | -31.2885
20.5592 | -35.2261 | -31.2948
20.5653 | -35.2264 | -31.3026
20.5714 | -35.2269 | -31.308
20.5775 | -35.2274 | -31.3139
20.5836 | -35.2278 | -31.3204
20.5897 | -35.2282 | -31.3268
20.5958 | -35.2288 | -31.3337
20.6019 | -35.2292 | -31.34
20.608 | -35.2299 | -31.3467
20.6141 | -35.2301 | -31.3525
20.6202 | -35.2305 | -31.3598
20.6263 | -35.231 | -31.3662
20.6324 | -35.2315 | -31.3726
20.6385 | -35.2321 | -31.379
20.6446 | -35.2326 | -31.3869
20.6507 | -35.2331 | -31.3942
20.6568 | -35.2336 | -31.4007
20.6629 | -35.234 | -31.4078
20.669 | -35.2344 | -31.4148
20.6751 | -35.2348 | -31.4217
20.6812 | -35.2356 | -31.4288
20.6873 | -35.236 | -31.4361
20.6934 | -35.2366 | -31.4418
20.6995 | -35.237 | -31.448
20.7057 | -35.2377 | -31.4542
20.7118 | -35.2383 | -31.4606
20.7179 | -35.2388 | -31.4664
20.724 | -35.2393 | -31.4724
20.7301 | -35.2399 | -31.4789
20.7362 | -35.2403 | -31.486
20.7423 | -35.2408 | -31.4927
20.7484 | -35.2414 | -31.4996
20.7545 | -35.2417 | -31.5071
20.7606 | -35.2422 | -31.5136
20.7667 | -35.2429 | -31.5206
20.7728 | -35.2433 | -31.5274
20.7789 | -35.2438 | -31.5344
20.785 | -35.2443 | -31.542
20.7911 | -35.2448 | -31.5482
20.7972 | -35.2454 | -31.5546
20.8033 | -35.2458 | -31.56
20.8094 | -35.2461 | -31.5664
20.8155 | -35.2467 | -31.5741
20.8216 | -35.2472 | -31.5806
20.8277 | -35.2479 | -31.5874
20.8338 | -35.2484 | -31.5939
20.8399 | -35.2488 | -31.5998
20.8461 | -35.2492 | -31.6065
20.8522 | -35.2496 | -31.6133
20.8583 | -35.2502 | -31.6197
20.8644 | -35.2506 | -31.6264
20.8705 | -35.251 | -31.632
20.8766 | -35.2514 | -31.6404
20.8827 | -35.2519 | -31.647
20.8888 | -35.2523 | -31.6534
20.8949 | -35.253 | -31.6599
20.901 | -35.2534 | -31.6677
20.9071 | -35.2539 | -31.6738
20.9132 | -35.2543 | -31.6798
20.9193 | -35.2547 | -31.6855
20.9254 | -35.2549 | -31.6921
20.9315 | -35.2555 | -31.6981
20.9376 | -35.256 | -31.705
20.9437 | -35.2566 | -31.7121
20.9498 | -35.2571 | -31.7192
20.9559 | -35.2577 | -31.7257
20.962 | -35.2582 | -31.7318
20.9681 | -35.2587 | -31.7373
20.9742 | -35.2592 | -31.7429
20.9803 | -35.2595 | -31.7493
20.9864 | -35.2601 | -31.7555
20.9926 | -35.2604 | -31.7619
20.9987 | -35.2607 | -31.7692
21.0048 | -35.2612 | -31.7756
21.0109 | -35.2615 | -31.7821
21.017 | -35.262 | -31.7887
21.0231 | -35.2627 | -31.7957
21.0292 | -35.2631 | -31.8014
21.0353 | -35.2638 | -31.8074
21.0414 | -35.2642 | -31.8133
21.0475 | -35.2646 | -31.821
21.0536 | -35.2648 | -31.8284
21.0597 | -35.2652 | -31.8353
21.0658 | -35.2656 | -31.8418
21.0719 | -35.2662 | -31.8496
21.078 | -35.2668 | -31.856
21.0841 | -35.2674 | -31.8638
21.0902 | -35.2676 | -31.8693
21.0963 | -35.268 | -31.8771
21.1024 | -35.2683 | -31.8818
21.1085 | -35.2686 | -31.8864
21.1146 | -35.2691 | -31.8924
21.1207 | -35.2697 | -31.8976
21.1268 | -35.2702 | -31.9047
21.133 | -35.2707 | -31.9099
21.1391 | -35.2711 | -31.9169
21.1452 | -35.2714 | -31.9231
21.1513 | -35.2719 | -31.93
21.1574 | -35.2724 | -31.9363
21.1635 | -35.2729 | -31.9428
21.1696 | -35.2735 | -31.9481
21.1757 | -35.2738 | -31.9524
21.1818 | -35.2744 | -31.958
21.1879 | -35.2749 | -31.9654
21.194 | -35.2753 | -31.9715
21.2001 | -35.2758 | -31.9787
21.2062 | -35.2764 | -31.9838
21.2123 | -35.2771 | -31.9911
21.2184 | -35.2775 | -31.997
21.2245 | -35.278 | -32.0025
21.2306 | -35.2786 | -32.0075
21.2367 | -35.2794 | -32.0145
21.2428 | -35.2799 | -32.0207
21.2489 | -35.2805 | -32.0267
21.255 | -35.2811 | -32.0344
21.2611 | -35.2816 | -32.0404
21.2672 | -35.282 | -32.0472
21.2733 | -35.2824 | -32.0521
21.2795 | -35.2828 | -32.0596
21.2856 | -35.2835 | -32.0667
21.2917 | -35.2838 | -32.0726
21.2978 | -35.2841 | -32.0777
21.3039 | -35.2844 | -32.0832
21.31 | -35.2847 | -32.0905
21.3161 | -35.285 | -32.0969
21.3222 | -35.2855 | -32.1039
21.3283 | -35.2861 | -32.1117
21.3344 | -35.2865 | -32.1182
21.3405 | -35.2872 | -32.1242
21.3466 | -35.2877 | -32.1303
21.3527 | -35.2882 | -32.1372
21.3588 | -35.2887 | -32.1424
21.3649 | -35.2892 | -32.1486
21.371 | -35.2897 | -32.1541
21.3771 | -35.29 | -32.1608
21.3832 | -35.2904 | -32.1664
21.3893 | -35.291 | -32.1717
21.3954 | -35.2915 | -32.1794
21.4015 | -35.2921 | -32.1839
21.4076 | -35.2925 | -32.1911
21.4137 | -35.293 | -32.1974
21.4199 | -35.2932 | -32.2042
21.426 | -35.2937 | -32.2099
21.4321 | -35.2941 | -32.2156
21.4382 | -35.2947 | -32.222
21.4443 | -35.2953 | -32.2281
21.4504 | -35.296 | -32.2342
21.4565 | -35.2963 | -32.2412
21.4626 | -35.2969 | -32.2474
21.4687 | -35.2975 | -32.2537
21.4748 | -35.2977 | -32.2596
21.4809 | -35.2983 | -32.2667
21.487 | -35.2988 | -32.2736
21.4931 | -35.2993 | -32.2808
21.4992 | -35.2998 | -32.2855
21.5053 | -35.3005 | -32.2921
21.5114 | -35.3009 | -32.2973
21.5175 | -35.3013 | -32.3034
21.5236 | -35.3013 | -32.3107
21.5297 | -35.3017 | -32.318
21.5358 | -35.3024 | -32.3253
21.5419 | -35.3028 | -32.3316
21.548 | -35.3031 | -32.3381
21.5541 | -35.3034 | -32.344
21.5602 | -35.3038 | -32.3505
21.5664 | -35.3037 | -32.3568
21.5725 | -35.3041 | -32.363
21.5786 | -35.3046 | -32.3704
21.5847 | -35.3053 | -32.3764
21.5908 | -35.3056 | -32.3826
21.5969 | -35.3062 | -32.3889
21.603 | -35.3066 | -32.3949
21.6091 | -35.3069 | -32.4008
21.6152 | -35.3074 | -32.4079
21.6213 | -35.3078 | -32.4138
21.6274 | -35.3082 | -32.4187
21.6335 | -35.3082 | -32.4234
21.6396 | -35.3086 | -32.4295
21.6457 | -35.3092 | -32.4354
21.6518 | -35.3098 | -32.4418
21.6579 | -35.3104 | -32.4477
21.664 | -35.3109 | -32.4542
21.6701 | -35.3114 | -32.4593
21.6762 | -35.3118 | -32.4658
21.6823 | -35.3121 | -32.4721
21.6884 | -35.3128 | -32.4773
21.6945 | -35.3131 | -32.4831
21.7006 | -35.3136 | -32.4897
21.7068 | -35.3141 | -32.4971
21.7129 | -35.3146 | -32.5036
21.719 | -35.315 | -32.5086
21.7251 | -35.3155 | -32.5139
21.7312 | -35.316 | -32.5197
21.7373 | -35.3162 | -32.5259
21.7434 | -35.317 | -32.5319
21.7495 | -35.3174 | -32.5385
21.7556 | -35.3179 | -32.5449
21.7617 | -35.3182 | -32.5517
21.7678 | -35.3186 | -32.5572
21.7739 | -35.3191 | -32.5647
21.78 | -35.3195 | -32.5709
21.7861 | -35.3201 | -32.5779
21.7922 | -35.3205 | -32.5841
21.7983 | -35.321 | -32.5895
21.8044 | -35.3214 | -32.5962
21.8105 | -35.3217 | -32.6014
21.8166 | -35.3221 | -32.6084
21.8227 | -35.3225 | -32.6147
21.8288 | -35.323 | -32.6204
21.8349 | -35.3235 | -32.6261
21.841 | -35.3242 | -32.633
21.8471 | -35.3247 | -32.6388
21.8533 | -35.3253 | -32.6442
21.8594 | -35.3258 | -32.6524
21.8655 | -35.3262 | -32.6585
21.8716 | -35.3265 | -32.6651
21.8777 | -35.3269 | -32.6713
21.8838 | -35.3272 | -32.6781
21.8899 | -35.3276 | -32.683
21.896 | -35.328 | -32.6886
21.9021 | -35.3285 | -32.6954
21.9082 | -35.329 | -32.7021
21.9143 | -35.3291 | -32.709
21.9204 | -35.3296 | -32.7164
21.9265 | -35.3299 | -32.7227
21.9326 | -35.3304 | -32.7277
21.9387 | -35.3309 | -32.7337
21.9448 | -35.3312 | -32.7398
21.9509 | -35.332 | -32.7458
21.957 | -35.3325 | -32.7526
21.9631 | -35.333 | -32.7591
21.9692 | -35.3333 | -32.7643
21.9753 | -35.3339 | -32.7699
21.9814 | -35.3344 | -32.7759
21.9875 | -35.3349 | -32.7806
21.9937 | -35.3353 | -32.7876
21.9998 | -35.3357 | -32.794
22.0059 | -35.3361 | -32.8003
22.012 | -35.3367 | -32.8063
22.0181 | -35.3372 | -32.8121
22.0242 | -35.3378 | -32.8181
22.0303 | -35.3382 | -32.824
22.0364 | -35.3387 | -32.831
22.0425 | -35.3391 | -32.8385
22.0486 | -35.3394 | -32.844
22.0547 | -35.3397 | -32.8504
22.0608 | -35.3402 | -32.8576
22.0669 | -35.3408 | -32.8639
22.073 | -35.3413 | -32.8695
22.0791 | -35.3418 | -32.8758
22.0852 | -35.3423 | -32.8814
22.0913 | -35.3427 | -32.886
22.0974 | -35.3431 | -32.8914
22.1035 | -35.3436 | -32.898
22.1096 | -35.3439 | -32.9034
22.1157 | -35.3444 | -32.9099
22.1218 | -35.3449 | -32.9163
22.1279 | -35.3454 | -32.9215
22.134 | -35.346 | -32.9271
22.1402 | -35.3465 | -32.9332
22.1463 | -35.347 | -32.9397
22.1524 | -35.3473 | -32.9453
22.1585 | -35.3476 | -32.9518
22.1646 | -35.3478 | -32.9575
22.1707 | -35.3482 | -32.9643
22.1768 | -35.3486 | -32.9715
22.1829 | -35.3491 | -32.9775
22.189 | -35.3493 | -32.9838
22.1951 | -35.3499 | -32.9891
22.2012 | -35.3504 | -32.9949
22.2073 | -35.3508 | -33.0018
22.2134 | -35.3512 | -33.0079
22.2195 | -35.3516 | -33.0136
22.2256 | -35.352 | -33.0193
22.2317 | -35.3525 | -33.0249
22.2378 | -35.3529 | -33.0316
22.2439 | -35.3534 | -33.0367
22.25 | -35.354 | -33.0425
22.2561 | -35.3546 | -33.0483
22.2622 | -35.3551 | -33.0541
22.2683 | -35.3555 | -33.0593
22.2744 | -35.3559 | -33.0653
22.2806 | -35.3564 | -33.0719
22.2867 | -35.3567 | -33.0778
22.2928 | -35.3572 | -33.0841
22.2989 | -35.3576 | -33.0906
22.305 | -35.358 | -33.0973
22.3111 | -35.3584 | -33.1033
22.3172 | -35.3588 | -33.1103
22.3233 | -35.3593 | -33.1163
22.3294 | -35.3597 | -33.1214
22.3355 | -35.3602 | -33.1274
22.3416 | -35.3606 | -33.1339
22.3477 | -35.361 | -33.1394
22.3538 | -35.3614 | -33.1453
22.3599 | -35.3617 | -33.1512
22.366 | -35.3621 | -33.1569
22.3721 | -35.3623 | -33.164
22.3782 | -35.3629 | -33.1693
22.3843 | -35.3634 | -33.1755
22.3904 | -35.3638 | -33.1825
22.3965 | -35.3642 | -33.1891
22.4026 | -35.3647 | -33.1942
22.4087 | -35.365 | -33.1999
22.4148 | -35.3655 | -33.2072
22.4209 | -35.3659 | -33.2132
22.4271 | -35.3662 | -33.2191
22.4332 | -35.3667 | -33.2256
22.4393 | -35.3673 | -33.2317
22.4454 | -35.3677 | -33.2367
22.4515 | -35.368 | -33.2426
22.4576 | -35.3685 | -33.2484
22.4637 | -35.369 | -33.2544
22.4698 | -35.3695 | -33.2595
22.4759 | -35.3699 | -33.2657
22.482 | -35.3703 | -33.2721
22.4881 | -35.3708 | -33.2782
22.4942 | -35.3711 | -33.2834
22.5003 | -35.3715 | -33.2899
22.5064 | -35.3718 | -33.2958
22.5125 | -35.3721 | -33.3006
22.5186 | -35.3726 | -33.3069
22.5247 | -35.3729 | -33.3134
22.5308 | -35.3734 | -33.319
22.5369 | -35.374 | -33.326
22.543 | -35.3744 | -33.3314
22.5491 | -35.375 | -33.337
22.5552 | -35.3753 | -33.3427
22.5613 | -35.3757 | -33.3498
22.5675 | -35.3762 | -33.3569
22.5736 | -35.3766 | -33.3625
22.5797 | -35.3771 | -33.3681
22.5858 | -35.3776 | -33.3744
22.5919 | -35.378 | -33.3801
22.598 | -35.3785 | -33.3863
22.6041 | -35.3788 | -33.3927
22.6102 | -35.3792 | -33.3976
22.6163 | -35.3801 | -33.4035
22.6224 | -35.3804 | -33.41
22.6285 | -35.3809 | -33.4166
22.6346 | -35.3812 | -33.4218
22.6407 | -35.3817 | -33.4277
22.6468 | -35.3822 | -33.4344
22.6529 | -35.3825 | -33.4395
22.659 | -35.383 | -33.4449
22.6651 | -35.3834 | -33.4511
22.6712 | -35.3839 | -33.4573
22.6773 | -35.3843 | -33.4633
22.6834 | -35.3848 | -33.4699
22.6895 | -35.3852 | -33.4756
22.6956 | -35.3855 | -33.4813
22.7017 | -35.386 | -33.4886
22.7079 | -35.3863 | -33.4956
22.714 | -35.3867 | -33.5018
22.7201 | -35.3871 | -33.5083
22.7262 | -35.3874 | -33.5132
22.7323 | -35.3879 | -33.519
22.7384 | -35.3884 | -33.5245
22.7445 | -35.3888 | -33.5292
22.7506 | -35.3893 | -33.5356
22.7567 | -35.3898 | -33.5416
22.7628 | -35.3903 | -33.5478
22.7689 | -35.3908 | -33.5532
22.775 | -35.3911 | -33.5594
22.7811 | -35.3916 | -33.5654
22.7872 | -35.3921 | -33.5698
22.7933 | -35.3926 | -33.576
22.7994 | -35.393 | -33.5823
22.8055 | -35.3935 | -33.5887
22.8116 | -35.3938 | -33.594
22.8177 | -35.3942 | -33.6
22.8238 | -35.3946 | -33.6069
22.8299 | -35.3953 | -33.6126
22.836 | -35.3957 | -33.6194
22.8421 | -35.3962 | -33.6256
22.8482 | -35.3967 | -33.6312
22.8544 | -35.3972 | -33.6379
22.8605 | -35.3976 | -33.6434
22.8666 | -35.398 | -33.6498
22.8727 | -35.3983 | -33.6557
22.8788 | -35.3986 | -33.6606
22.8849 | -35.3991 | -33.667
22.891 | -35.3995 | -33.6732
22.8971 | -35.4 | -33.6788
22.9032 | -35.4005 | -33.6853
22.9093 | -35.4009 | -33.6909
22.9154 | -35.4013 | -33.6963
22.9215 | -35.4017 | -33.7023
22.9276 | -35.4021 | -33.7088
22.9337 | -35.4026 | -33.7161
22.9398 | -35.403 | -33.7223
22.9459 | -35.4034 | -33.7284
22.952 | -35.4037 | -33.7353
22.9581 | -35.4042 | -33.741
22.9642 | -35.4047 | -33.7471
22.9703 | -35.4051 | -33.7538
22.9764 | -35.4056 | -33.7601
22.9825 | -35.406 | -33.7661
22.9886 | -35.4064 | -33.7722
22.9948 | -35.4069 | -33.7785
23.0009 | -35.4072 | -33.7847
23.007 | -35.4078 | -33.7901
23.0131 | -35.4082 | -33.7962
23.0192 | -35.4085 | -33.8013
23.0253 | -35.4092 | -33.8057
23.0314 | -35.4097 | -33.8117
23.0375 | -35.4102 | -33.8185
23.0436 | -35.4107 | -33.8244
23.0497 | -35.4109 | -33.831
23.0558 | -35.4113 | -33.8362
23.0619 | -35.4119 | -33.8419
23.068 | -35.4122 | -33.8476
23.0741 | -35.4125 | -33.8532
23.0802 | -35.413 | -33.8595
23.0863 | -35.4134 | -33.8655
23.0924 | -35.4138 | -33.8716
23.0985 | -35.4142 | -33.8775
23.1046 | -35.4147 | -33.8835
23.1107 | -35.415 | -33.8897
23.1168 | -35.4154 | -33.8959
23.1229 | -35.4159 | -33.9012
23.129 | -35.4163 | -33.9071
23.1351 | -35.4168 | -33.9126
23.1413 | -35.4171 | -33.9198
23.1474 | -35.4175 | -33.9248
23.1535 | -35.418 | -33.931
23.1596 | -35.4184 | -33.9371
23.1657 | -35.4188 | -33.9421
23.1718 | -35.4193 | -33.9475
23.1779 | -35.4196 | -33.954
23.184 | -35.4203 | -33.9587
23.1901 | -35.4206 | -33.9648
23.1962 | -35.421 | -33.9712
23.2023 | -35.4214 | -33.9775
23.2084 | -35.4218 | -33.9832
23.2145 | -35.422 | -33.9909
23.2206 | -35.4224 | -33.9964
23.2267 | -35.4228 | -34.0018
23.2328 | -35.4232 | -34.0078
23.2389 | -35.4235 | -34.0147
23.245 | -35.4241 | -34.0203
23.2511 | -35.4247 | -34.0257
23.2572 | -35.425 | -34.0305
23.2633 | -35.4253 | -34.0363
23.2694 | -35.4258 | -34.0419
23.2755 | -35.4262 | -34.0486
23.2817 | -35.4267 | -34.0543
23.2878 | -35.4272 | -34.0607
23.2939 | -35.4274 | -34.0677
23.3 | -35.4277 | -34.0743
23.3061 | -35.4283 | -34.0805
23.3122 | -35.4288 | -34.0863
23.3183 | -35.4291 | -34.0921
23.3244 | -35.4294 | -34.0969
23.3305 | -35.4298 | -34.1033
23.3366 | -35.4302 | -34.1099
23.3427 | -35.4307 | -34.1154
23.3488 | -35.4311 | -34.1214
23.3549 | -35.4316 | -34.1286
23.361 | -35.432 | -34.135
23.3671 | -35.4324 | -34.141
23.3732 | -35.4328 | -34.1466
23.3793 | -35.4332 | -34.1522
23.3854 | -35.4335 | -34.1577
23.3915 | -35.434 | -34.1642
23.3976 | -35.4345 | -34.17
23.4037 | -35.4349 | -34.1757
23.4098 | -35.4353 | -34.1816
23.4159 | -35.4357 | -34.1876
23.422 | -35.4363 | -34.1935
23.4282 | -35.4364 | -34.1999
23.4343 | -35.4368 | -34.2058
23.4404 | -35.4371 | -34.2125
23.4465 | -35.4374 | -34.2184
23.4526 | -35.4376 | -34.2244
23.4587 | -35.438 | -34.2301
23.4648 | -35.4382 | -34.2364
23.4709 | -35.4386 | -34.2429
23.477 | -35.439 | -34.2484
23.4831 | -35.4394 | -34.2544
23.4892 | -35.4399 | -34.2601
23.4953 | -35.4403 | -34.2666
23.5014 | -35.4406 | -34.2715
23.5075 | -35.441 | -34.2771
23.5136 | -35.4415 | -34.2835
23.5197 | -35.4419 | -34.2893
23.5258 | -35.4423 | -34.296
23.5319 | -35.4428 | -34.3026
23.538 | -35.4431 | -34.3081
23.5441 | -35.4435 | -34.3137
23.5502 | -35.4438 | -34.3201
23.5563 | -35.4443 | -34.3265
23.5624 | -35.4446 | -34.3328
23.5686 | -35.445 | -34.3382
23.5747 | -35.4454 | -34.3447
23.5808 | -35.4458 | -34.3508
23.5869 | -35.4462 | -34.3556
23.593 | -35.4466 | -34.3623
23.5991 | -35.447 | -34.3688
23.6052 | -35.4474 | -34.3735
23.6113 | -35.4477 | -34.3798
23.6174 | -35.4481 | -34.3854
23.6235 | -35.4485 | -34.3914
23.6296 | -35.4489 | -34.3972
23.6357 | -35.4493 | -34.4034
23.6418 | -35.4497 | -34.4096
23.6479 | -35.45 | -34.4158
23.654 | -35.4504 | -34.4223
23.6601 | -35.4508 | -34.4284
23.6662 | -35.4512 | -34.4349
23.6723 | -35.4517 | -34.4417
23.6784 | -35.4522 | -34.4479
23.6845 | -35.4524 | -34.4539
23.6906 | -35.4528 | -34.4596
23.6967 | -35.4532 | -34.4659
23.7028 | -35.4536 | -34.4718
23.7089 | -35.4538 | -34.4781
23.7151 | -35.4542 | -34.4837
23.7212 | -35.4546 | -34.4897
23.7273 | -35.4549 | -34.4954
23.7334 | -35.4553 | -34.5012
23.7395 | -35.4557 | -34.5071
23.7456 | -35.456 | -34.5124
23.7517 | -35.4564 | -34.5186
23.7578 | -35.4569 | -34.5246
23.7639 | -35.4573 | -34.5308
23.77 | -35.4577 | -34.5374
23.7761 | -35.4581 | -34.5431
23.7822 | -35.4583 | -34.5481
23.7883 | -35.4586 | -34.5553
23.7944 | -35.459 | -34.5613
23.8005 | -35.4594 | -34.5678
23.8066 | -35.4598 | -34.5738
23.8127 | -35.4603 | -34.5799
23.8188 | -35.4607 | -34.5855
23.8249 | -35.4611 | -34.5909
23.831 | -35.4615 | -34.5972
23.8371 | -35.4619 | -34.6033
23.8432 | -35.4624 | -34.6095
23.8493 | -35.4628 | -34.6143
23.8555 | -35.4631 | -34.6204
23.8616 | -35.4634 | -34.6267
23.8677 | -35.4638 | -34.6323
23.8738 | -35.4642 | -34.6393
23.8799 | -35.4646 | -34.646
23.886 | -35.465 | -34.6528
23.8921 | -35.4654 | -34.6587
23.8982 | -35.4658 | -34.6652
23.9043 | -35.4661 | -34.6718
23.9104 | -35.4665 | -34.6783
23.9165 | -35.4668 | -34.6841
23.9226 | -35.4673 | -34.6898
23.9287 | -35.4678 | -34.6963
23.9348 | -35.468 | -34.7019
23.9409 | -35.4685 | -34.7075
23.947 | -35.4689 | -34.7139
23.9531 | -35.4693 | -34.7199
23.9592 | -35.4697 | -34.7256
23.9653 | -35.4703 | -34.7314
23.9714 | -35.4708 | -34.737
23.9775 | -35.4713 | -34.7424
23.9836 | -35.4717 | -34.7488
23.9897 | -35.4722 | -34.7558
23.9958 | -35.4727 | -34.7623
24.002 | -35.4732 | -34.7679
24.0081 | -35.4737 | -34.7736
24.0142 | -35.474 | -34.7795
24.0203 | -35.4744 | -34.7853
24.0264 | -35.4748 | -34.7911
24.0325 | -35.4752 | -34.7967
24.0386 | -35.4754 | -34.8034
24.0447 | -35.4759 | -34.8094
24.0508 | -35.4762 | -34.8155
24.0569 | -35.4767 | -34.8219
24.063 | -35.4771 | -34.8276
24.0691 | -35.4776 | -34.8332
24.0752 | -35.478 | -34.8387
24.0813 | -35.4784 | -34.8445
24.0874 | -35.4789 | -34.8505
24.0935 | -35.4792 | -34.8561
24.0996 | -35.4795 | -34.8623
24.1057 | -35.4799 | -34.8682
24.1118 | -35.4802 | -34.874
24.1179 | -35.4805 | -34.8804
24.124 | -35.4809 | -34.8857
24.1301 | -35.4813 | -34.8916
24.1362 | -35.4817 | -34.8969
24.1424 | -35.482 | -34.9033
24.1485 | -35.4825 | -34.9093
24.1546 | -35.4829 | -34.9154
24.1607 | -35.4834 | -34.9216
24.1668 | -35.4837 | -34.9273
24.1729 | -35.4842 | -34.9338
24.179 | -35.4846 | -34.9398
24.1851 | -35.485 | -34.9463
24.1912 | -35.4854 | -34.9525
24.1973 | -35.4858 | -34.9586
24.2034 | -35.4862 | -34.9649
24.2095 | -35.4866 | -34.9715
24.2156 | -35.4869 | -34.9775
24.2217 | -35.4873 | -34.9834
24.2278 | -35.4876 | -34.9888
24.2339 | -35.488 | -34.9948
24.24 | -35.4884 | -35.0013
24.2461 | -35.4888 | -35.0077
24.2522 | -35.4891 | -35.0134
24.2583 | -35.4896 | -35.0204
24.2644 | -35.4901 | -35.027
24.2705 | -35.4906 | -35.0329
24.2766 | -35.4909 | -35.0385
24.2827 | -35.4913 | -35.0437
24.2889 | -35.4918 | -35.0496
24.295 | -35.4923 | -35.0559
24.3011 | -35.4925 | -35.0619
24.3072 | -35.4929 | -35.067
24.3133 | -35.4933 | -35.0729
24.3194 | -35.4936 | -35.079
24.3255 | -35.494 | -35.0859
24.3316 | -35.4943 | -35.0918
24.3377 | -35.4946 | -35.0984
24.3438 | -35.495 | -35.1043
24.3499 | -35.4953 | -35.1105
24.356 | -35.4957 | -35.1167
24.3621 | -35.496 | -35.123
24.3682 | -35.4965 | -35.1289
24.3743 | -35.4969 | -35.1354
24.3804 | -35.4973 | -35.1409
24.3865 | -35.4976 | -35.1463
24.3926 | -35.498 | -35.1526
24.3987 | -35.4983 | -35.1591
24.4048 | -35.4988 | -35.165
24.4109 | -35.4992 | -35.1706
24.417 | -35.4997 | -35.1769
24.4231 | -35.5001 | -35.1834
24.4293 | -35.5006 | -35.1888
24.4354 | -35.501 | -35.195
24.4415 | -35.5015 | -35.2014
24.4476 | -35.5019 | -35.2068
24.4537 | -35.5024 | -35.213
24.4598 | -35.5027 | -35.2194
24.4659 | -35.5031 | -35.2255
24.472 | -35.5034 | -35.2317
24.4781 | -35.5038 | -35.2378
24.4842 | -35.504 | -35.2441
24.4903 | -35.5044 | -35.2496
24.4964 | -35.5047 | -35.2558
24.5025 | -35.5051 | -35.2619
24.5086 | -35.5054 | -35.2683
24.5147 | -35.5058 | -35.2744
24.5208 | -35.5062 | -35.28
24.5269 | -35.5067 | -35.2857
24.533 | -35.5071 | -35.2916
24.5391 | -35.5075 | -35.2972
24.5452 | -35.5079 | -35.3037
24.5513 | -35.5083 | -35.3105
24.5574 | -35.5087 | -35.3164
24.5635 | -35.5091 | -35.3224
24.5696 | -35.5096 | -35.3283
24.5758 | -35.5099 | -35.3348
24.5819 | -35.5103 | -35.3408
24.588 | -35.5109 | -35.3458
24.5941 | -35.5113 | -35.3521
24.6002 | -35.5117 | -35.3584
24.6063 | -35.512 | -35.3638
24.6124 | -35.5123 | -35.3702
24.6185 | -35.5127 | -35.3764
24.6246 | -35.513 | -35.382
24.6307 | -35.5134 | -35.3884
24.6368 | -35.5137 | -35.3949
24.6429 | -35.514 | -35.4009
24.649 | -35.5145 | -35.4066
24.6551 | -35.5149 | -35.4126
24.6612 | -35.5152 | -35.4186
24.6673 | -35.5157 | -35.4248
24.6734 | -35.5159 | -35.431
24.6795 | -35.5163 | -35.4369
24.6856 | -35.5167 | -35.4436
24.6917 | -35.5171 | -35.4499
24.6978 | -35.5175 | -35.4553
24.7039 | -35.5178 | -35.4614
24.71 | -35.5182 | -35.4672
24.7162 | -35.5186 | -35.4736
24.7223 | -35.519 | -35.4796
24.7284 | -35.5194 | -35.4857
24.7345 | -35.5198 | -35.4917
24.7406 | -35.5202 | -35.4977
24.7467 | -35.5206 | -35.5041
24.7528 | -35.521 | -35.5098
24.7589 | -35.5213 | -35.5153
24.765 | -35.5217 | -35.5213
24.7711 | -35.5221 | -35.5272
24.7772 | -35.5225 | -35.5332
24.7833 | -35.5229 | -35.5391
24.7894 | -35.5233 | -35.5449
24.7955 | -35.5236 | -35.5511
24.8016 | -35.5239 | -35.557
24.8077 | -35.5243 | -35.5629
24.8138 | -35.5247 | -35.5694
24.8199 | -35.5251 | -35.5751
24.826 | -35.5255 | -35.5809
24.8321 | -35.5259 | -35.5875
24.8382 | -35.5262 | -35.5938
24.8443 | -35.5265 | -35.5997
24.8504 | -35.5269 | -35.606
24.8565 | -35.5272 | -35.6122
24.8627 | -35.5277 | -35.6189
24.8688 | -35.5281 | -35.6251
24.8749 | -35.5286 | -35.6311
24.881 | -35.5289 | -35.6377
24.8871 | -35.5294 | -35.6438
24.8932 | -35.5298 | -35.65
24.8993 | -35.5302 | -35.6562
24.9054 | -35.5307 | -35.6623
24.9115 | -35.5311 | -35.6683
24.9176 | -35.5315 | -35.6745
24.9237 | -35.5319 | -35.6807
24.9298 | -35.5322 | -35.6867
24.9359 | -35.5326 | -35.6926
24.942 | -35.5329 | -35.6987
24.9481 | -35.5333 | -35.7047
24.9542 | -35.5337 | -35.711
24.9603 | -35.534 | -35.7171
24.9664 | -35.5344 | -35.7228
24.9725 | -35.5348 | -35.729
24.9786 | -35.5352 | -35.7348
24.9847 | -35.5356 | -35.7409
24.9908 | -35.536 | -35.7472
24.9969 | -35.5363 | -35.7529
25.0031 | -35.5367 | -35.7591
25.0092 | -35.5371 | -35.7652
25.0153 | -35.5374 | -35.7706
25.0214 | -35.5378 | -35.7767
25.0275 | -35.5382 | -35.7826
25.0336 | -35.5386 | -35.7885
25.0397 | -35.539 | -35.795
25.0458 | -35.5394 | -35.8011
25.0519 | -35.5398 | -35.8074
25.058 | -35.5401 | -35.8133
25.0641 | -35.5405 | -35.8195
25.0702 | -35.5409 | -35.8257
25.0763 | -35.5413 | -35.8315
25.0824 | -35.5416 | -35.8378
25.0885 | -35.542 | -35.8435
25.0946 | -35.5425 | -35.8492
25.1007 | -35.5429 | -35.8551
25.1068 | -35.5433 | -35.8611
25.1129 | -35.5436 | -35.8668
25.119 | -35.544 | -35.8726
25.1251 | -35.5444 | -35.8788
25.1312 | -35.5448 | -35.885
25.1373 | -35.5451 | -35.8914
25.1435 | -35.5455 | -35.8975
25.1496 | -35.5459 | -35.9036
25.1557 | -35.5463 | -35.9092
25.1618 | -35.5466 | -35.9156
25.1679 | -35.547 | -35.9213
25.174 | -35.5474 | -35.9272
25.1801 | -35.5478 | -35.9336
25.1862 | -35.5481 | -35.9396
25.1923 | -35.5485 | -35.9458
25.1984 | -35.549 | -35.9521
25.2045 | -35.5493 | -35.9579
25.2106 | -35.5498 | -35.9639
25.2167 | -35.5502 | -35.9702
25.2228 | -35.5506 | -35.9758
25.2289 | -35.5509 | -35.9817
25.235 | -35.5513 | -35.9875
25.2411 | -35.5517 | -35.9939
25.2472 | -35.5521 | -36.0002
25.2533 | -35.5525 | -36.0063
25.2594 | -35.5529 | -36.0127
25.2655 | -35.5533 | -36.0186
25.2716 | -35.5537 | -36.0246
25.2777 | -35.5541 | -36.0309
25.2838 | -35.5545 | -36.0374
25.29 | -35.5548 | -36.0431
25.2961 | -35.5552 | -36.0491
25.3022 | -35.5556 | -36.0549
25.3083 | -35.5559 | -36.0608
25.3144 | -35.5563 | -36.0671
25.3205 | -35.5567 | -36.0729
25.3266 | -35.557 | -36.0788
25.3327 | -35.5574 | -36.0846
25.3388 | -35.5578 | -36.0907
25.3449 | -35.5581 | -36.0969
25.351 | -35.5585 | -36.103
25.3571 | -35.5589 | -36.1087
25.3632 | -35.5593 | -36.1148
25.3693 | -35.5596 | -36.1208
25.3754 | -35.56 | -36.1272
25.3815 | -35.5604 | -36.133
25.3876 | -35.5608 | -36.1391
25.3937 | -35.5611 | -36.1452
25.3998 | -35.5614 | -36.1512
25.4059 | -35.5618 | -36.1572
25.412 | -35.5622 | -36.1634
25.4181 | -35.5626 | -36.1696
25.4242 | -35.563 | -36.1757
25.4304 | -35.5634 | -36.1818
25.4365 | -35.5638 | -36.1881
25.4426 | -35.5642 | -36.1938
25.4487 | -35.5646 | -36.2
25.4548 | -35.565 | -36.2058
25.4609 | -35.5654 | -36.2121
25.467 | -35.5658 | -36.2183
25.4731 | -35.5662 | -36.2243
25.4792 | -35.5665 | -36.2302
25.4853 | -35.5669 | -36.2358
25.4914 | -35.5673 | -36.2418
25.4975 | -35.5677 | -36.2478
25.5036 | -35.5682 | -36.2537
25.5097 | -35.5685 | -36.2593
25.5158 | -35.5689 | -36.2651
25.5219 | -35.5693 | -36.2709
25.528 | -35.5697 | -36.2768
25.5341 | -35.5701 | -36.2825
25.5402 | -35.5705 | -36.2883
25.5463 | -35.5709 | -36.2947
25.5524 | -35.5713 | -36.3009
25.5585 | -35.5716 | -36.307
25.5646 | -35.572 | -36.313
25.5707 | -35.5723 | -36.3187
25.5769 | -35.5727 | -36.3246
25.583 | -35.5731 | -36.3308
25.5891 | -35.5735 | -36.3371
25.5952 | -35.5738 | -36.3436
25.6013 | -35.5743 | -36.3498
25.6074 | -35.5746 | -36.3559
25.6135 | -35.575 | -36.3618
25.6196 | -35.5754 | -36.3682
25.6257 | -35.5758 | -36.3744
25.6318 | -35.5761 | -36.3801
25.6379 | -35.5765 | -36.386
25.644 | -35.5769 | -36.3912
25.6501 | -35.5773 | -36.3966
25.6562 | -35.5777 | -36.4026
25.6623 | -35.5781 | -36.4085
25.6684 | -35.5784 | -36.4147
25.6745 | -35.5788 | -36.421
25.6806 | -35.5792 | -36.4268
25.6867 | -35.5796 | -36.4325
25.6928 | -35.5799 | -36.4383
25.6989 | -35.5803 | -36.4442
25.705 | -35.5806 | -36.4505
25.7111 | -35.581 | -36.4567
25.7173 | -35.5814 | -36.4625
25.7234 | -35.5817 | -36.4687
25.7295 | -35.5821 | -36.4747
25.7356 | -35.5824 | -36.4806
25.7417 | -35.5828 | -36.4867
25.7478 | -35.5832 | -36.4922
25.7539 | -35.5835 | -36.4979
25.76 | -35.5839 | -36.5038
25.7661 | -35.5843 | -36.5097
25.7722 | -35.5847 | -36.5152
25.7783 | -35.5851 | -36.5212
25.7844 | -35.5855 | -36.5269
25.7905 | -35.5859 | -36.5331
25.7966 | -35.5862 | -36.539
25.8027 | -35.5865 | -36.545
25.8088 | -35.5869 | -36.5512
25.8149 | -35.5873 | -36.5572
25.821 | -35.5876 | -36.5635
25.8271 | -35.588 | -36.5695
25.8332 | -35.5884 | -36.5756
25.8393 | -35.5889 | -36.5816
25.8454 | -35.5893 | -36.5877
25.8515 | -35.5897 | -36.5938
25.8576 | -35.5901 | -36.5998
25.8638 | -35.5905 | -36.606
25.8699 | -35.5908 | -36.6123
25.876 | -35.5912 | -36.6183
25.8821 | -35.5916 | -36.6245
25.8882 | -35.592 | -36.6304
25.8943 | -35.5924 | -36.6363
25.9004 | -35.5927 | -36.6422
25.9065 | -35.5931 | -36.6481
25.9126 | -35.5935 | -36.6541
25.9187 | -35.5939 | -36.66
25.9248 | -35.5942 | -36.6663
25.9309 | -35.5946 | -36.6725
25.937 | -35.595 | -36.6786
25.9431 | -35.5954 | -36.6845
25.9492 | -35.5958 | -36.6903
25.9553 | -35.5961 | -36.6961
25.9614 | -35.5965 | -36.7017
25.9675 | -35.5969 | -36.7074
25.9736 | -35.5973 | -36.7133
25.9797 | -35.5977 | -36.7196
25.9858 | -35.598 | -36.7254
25.9919 | -35.5983 | -36.7313
25.998 | -35.5987 | -36.7369
26.0042 | -35.5991 | -36.7429
26.0103 | -35.5995 | -36.7489
26.0164 | -35.5999 | -36.7549
26.0225 | -35.6002 | -36.7609
26.0286 | -35.6006 | -36.7668
26.0347 | -35.601 | -36.7723
26.0408 | -35.6014 | -36.7788
26.0469 | -35.6018 | -36.7845
26.053 | -35.6021 | -36.7908
26.0591 | -35.6025 | -36.7968
26.0652 | -35.6029 | -36.8026
26.0713 | -35.6033 | -36.8086
26.0774 | -35.6037 | -36.8142
26.0835 | -35.6041 | -36.8203
26.0896 | -35.6046 | -36.8262
26.0957 | -35.605 | -36.8319
26.1018 | -35.6054 | -36.8374
26.1079 | -35.6058 | -36.8432
26.114 | -35.6062 | -36.8493
26.1201 | -35.6065 | -36.8553
26.1262 | -35.6068 | -36.8615
26.1323 | -35.6072 | -36.8672
26.1384 | -35.6077 | -36.8729
26.1445 | -35.6081 | -36.8788
26.1507 | -35.6085 | -36.8849
26.1568 | -35.6088 | -36.891
26.1629 | -35.6092 | -36.8968
26.169 | -35.6096 | -36.9026
26.1751 | -35.61 | -36.9083
26.1812 | -35.6104 | -36.9143
26.1873 | -35.6108 | -36.9202
26.1934 | -35.6112 | -36.926
26.1995 | -35.6116 | -36.9322
26.2056 | -35.612 | -36.9376
26.2117 | -35.6124 | -36.9433
26.2178 | -35.6128 | -36.9492
26.2239 | -35.6132 | -36.9549
26.23 | -35.6135 | -36.9611
26.2361 | -35.6139 | -36.9668
26.2422 | -35.6144 | -36.9726
26.2483 | -35.6147 | -36.9784
26.2544 | -35.6151 | -36.9843
26.2605 | -35.6154 | -36.9902
26.2666 | -35.6158 | -36.9959
26.2727 | -35.6161 | -37.0017
26.2788 | -35.6165 | -37.0074
26.2849 | -35.6169 | -37.0131
26.2911 | -35.6172 | -37.019
26.2972 | -35.6176 | -37.025
26.3033 | -35.618 | -37.0309
26.3094 | -35.6184 | -37.0368
26.3155 | -35.6188 | -37.0423
26.3216 | -35.6191 | -37.0483
26.3277 | -35.6195 | -37.0542
26.3338 | -35.6199 | -37.0597
26.3399 | -35.6203 | -37.0651
26.346 | -35.6207 | -37.0705
26.3521 | -35.621 | -37.0762
26.3582 | -35.6214 | -37.082
26.3643 | -35.6218 | -37.0875
26.3704 | -35.6222 | -37.0936
26.3765 | -35.6226 | -37.0996
26.3826 | -35.623 | -37.1056
26.3887 | -35.6233 | -37.1113
26.3948 | -35.6237 | -37.1171
26.4009 | -35.6241 | -37.1225
26.407 | -35.6245 | -37.1282
26.4131 | -35.6248 | -37.134
26.4192 | -35.6252 | -37.1399
26.4253 | -35.6256 | -37.1456
26.4314 | -35.626 | -37.1513
26.4376 | -35.6264 | -37.1569
26.4437 | -35.6267 | -37.163
26.4498 | -35.6271 | -37.1694
26.4559 | -35.6275 | -37.1748
26.462 | -35.6278 | -37.1807
26.4681 | -35.6282 | -37.1867
26.4742 | -35.6286 | -37.1926
26.4803 | -35.629 | -37.1983
26.4864 | -35.6294 | -37.2038
26.4925 | -35.6297 | -37.2096
26.4986 | -35.6301 | -37.2155
26.5047 | -35.6305 | -37.2214
26.5108 | -35.6309 | -37.227
26.5169 | -35.6313 | -37.2325
26.523 | -35.6316 | -37.2379
26.5291 | -35.632 | -37.2438
26.5352 | -35.6323 | -37.2498
26.5413 | -35.6327 | -37.2558
26.5474 | -35.6331 | -37.2616
26.5535 | -35.6336 | -37.2674
26.5596 | -35.6339 | -37.2732
26.5657 | -35.6343 | -37.2792
26.5718 | -35.6347 | -37.2851
26.578 | -35.635 | -37.2909
26.5841 | -35.6355 | -37.2967
26.5902 | -35.6358 | -37.3023
26.5963 | -35.6362 | -37.3078
26.6024 | -35.6367 | -37.3139
26.6085 | -35.6371 | -37.3197
26.6146 | -35.6374 | -37.3253
26.6207 | -35.6378 | -37.3309
26.6268 | -35.6382 | -37.3367
26.6329 | -35.6386 | -37.3424
26.639 | -35.639 | -37.3483
26.6451 | -35.6393 | -37.3546
26.6512 | -35.6397 | -37.3605
26.6573 | -35.6401 | -37.3663
26.6634 | -35.6404 | -37.3717
26.6695 | -35.6407 | -37.3775
26.6756 | -35.6411 | -37.3833
26.6817 | -35.6415 | -37.389
26.6878 | -35.6419 | -37.3949
26.6939 | -35.6422 | -37.4004
26.7 | -35.6426 | -37.4059
26.7061 | -35.643 | -37.4113
26.7122 | -35.6434 | -37.4171
26.7183 | -35.6438 | -37.4231
26.7245 | -35.6441 | -37.4291
26.7306 | -35.6445 | -37.4348
26.7367 | -35.6449 | -37.4401
26.7428 | -35.6453 | -37.4456
26.7489 | -35.6457 | -37.4511
26.755 | -35.6461 | -37.4563
26.7611 | -35.6465 | -37.4623
26.7672 | -35.6469 | -37.4681
26.7733 | -35.6473 | -37.4736
26.7794 | -35.6476 | -37.4792
26.7855 | -35.648 | -37.4853
26.7916 | -35.6483 | -37.4908
26.7977 | -35.6487 | -37.4967
26.8038 | -35.649 | -37.5024
26.8099 | -35.6494 | -37.5084
26.816 | -35.6497 | -37.5141
26.8221 | -35.6501 | -37.5195
26.8282 | -35.6504 | -37.5251
26.8343 | -35.6508 | -37.5307
26.8404 | -35.6511 | -37.5359
26.8465 | -35.6515 | -37.5417
26.8526 | -35.6519 | -37.5475
26.8587 | -35.6523 | -37.5535
26.8649 | -35.6527 | -37.5592
26.871 | -35.653 | -37.5649
26.8771 | -35.6533 | -37.5702
26.8832 | -35.6537 | -37.5754
26.8893 | -35.6541 | -37.5813
26.8954 | -35.6545 | -37.5871
26.9015 | -35.6549 | -37.593
26.9076 | -35.6553 | -37.599
26.9137 | -35.6557 | -37.6043
26.9198 | -35.6562 | -37.6103
26.9259 | -35.6565 | -37.6158
26.932 | -35.6568 | -37.6212
26.9381 | -35.6573 | -37.6266
26.9442 | -35.6576 | -37.6317
26.9503 | -35.658 | -37.6374
26.9564 | -35.6584 | -37.6429
26.9625 | -35.6588 | -37.6487
26.9686 | -35.6592 | -37.6541
26.9747 | -35.6596 | -37.6598
26.9808 | -35.66 | -37.665
26.9869 | -35.6603 | -37.6704
26.993 | -35.6607 | -37.6754
26.9991 | -35.661 | -37.6807
27.0052 | -35.6614 | -37.687
27.0114 | -35.6618 | -37.6925
27.0175 | -35.6621 | -37.6979
27.0236 | -35.6625 | -37.7032
27.0297 | -35.6629 | -37.7092
27.0358 | -35.6632 | -37.7144
27.0419 | -35.6635 | -37.7204
27.048 | -35.6639 | -37.7258
27.0541 | -35.6642 | -37.7313
27.0602 | -35.6646 | -37.7365
27.0663 | -35.665 | -37.7422
27.0724 | -35.6654 | -37.7481
27.0785 | -35.6657 | -37.7541
27.0846 | -35.6661 | -37.7593
27.0907 | -35.6665 | -37.7649
27.0968 | -35.6667 | -37.7705
27.1029 | -35.6671 | -37.7757
27.109 | -35.6675 | -37.7811
27.1151 | -35.6679 | -37.7865
27.1212 | -35.6682 | -37.7913
27.1273 | -35.6686 | -37.7974
27.1334 | -35.6691 | -37.803
27.1395 | -35.6695 | -37.8083
27.1456 | -35.6699 | -37.814
27.1518 | -35.6703 | -37.8192
27.1579 | -35.6707 | -37.8245
27.164 | -35.671 | -37.8301
27.1701 | -35.6714 | -37.8359
27.1762 | -35.6717 | -37.8413
27.1823 | -35.672 | -37.8471
27.1884 | -35.6724 | -37.8524
27.1945 | -35.6728 | -37.8579
27.2006 | -35.6731 | -37.8637
27.2067 | -35.6735 | -37.8693
27.2128 | -35.6738 | -37.8749
27.2189 | -35.6742 | -37.8805
27.225 | -35.6746 | -37.8863
27.2311 | -35.6749 | -37.8918
27.2372 | -35.6753 | -37.8977
27.2433 | -35.6757 | -37.9032
27.2494 | -35.676 | -37.909
27.2555 | -35.6764 | -37.9151
27.2616 | -35.6767 | -37.9204
27.2677 | -35.6771 | -37.9259
27.2738 | -35.6775 | -37.9311
27.2799 | -35.6779 | -37.9364
27.286 | -35.6783 | -37.9415
27.2921 | -35.6787 | -37.9473
27.2983 | -35.679 | -37.9528
27.3044 | -35.6793 | -37.9585
27.3105 | -35.6797 | -37.9639
27.3166 | -35.6799 | -37.9693
27.3227 | -35.6803 | -37.9748
27.3288 | -35.6807 | -37.9803
27.3349 | -35.6811 | -37.9854
27.341 | -35.6814 | -37.9913
27.3471 | -35.6817 | -37.9968
27.3532 | -35.6821 | -38.002
27.3593 | -35.6825 | -38.0076
27.3654 | -35.6829 | -38.0132
27.3715 | -35.6833 | -38.0183
27.3776 | -35.6837 | -38.024
27.3837 | -35.684 | -38.0295
27.3898 | -35.6843 | -38.0347
27.3959 | -35.6847 | -38.0402
27.402 | -35.6851 | -38.0457
27.4081 | -35.6854 | -38.0512
27.4142 | -35.6858 | -38.057
27.4203 | -35.6862 | -38.0626
27.4264 | -35.6865 | -38.0684
27.4325 | -35.6869 | -38.0743
27.4387 | -35.6873 | -38.0792
27.4448 | -35.6877 | -38.0847
27.4509 | -35.6881 | -38.09
27.457 | -35.6885 | -38.0954
27.4631 | -35.6889 | -38.1011
27.4692 | -35.6892 | -38.1065
27.4753 | -35.6896 | -38.1115
27.4814 | -35.6899 | -38.1169
27.4875 | -35.6902 | -38.1226
27.4936 | -35.6906 | -38.1279
27.4997 | -35.691 | -38.1332
27.5058 | -35.6913 | -38.1387
27.5119 | -35.6916 | -38.1436
27.518 | -35.692 | -38.1494
27.5241 | -35.6923 | -38.1546
27.5302 | -35.6927 | -38.16
27.5363 | -35.693 | -38.1658
27.5424 | -35.6934 | -38.1706
27.5485 | -35.6938 | -38.1756
27.5546 | -35.6941 | -38.1809
27.5607 | -35.6945 | -38.1862
27.5668 | -35.6948 | -38.191
27.5729 | -35.6952 | -38.1961
27.5791 | -35.6955 | -38.2011
27.5852 | -35.6959 | -38.2066
27.5913 | -35.6963 | -38.2122
27.5974 | -35.6967 | -38.2178
27.6035 | -35.6971 | -38.2233
27.6096 | -35.6974 | -38.2286
27.6157 | -35.6978 | -38.2337
27.6218 | -35.6981 | -38.2393
27.6279 | -35.6984 | -38.2449
27.634 | -35.6987 | -38.2505
27.6401 | -35.6991 | -38.256
27.6462 | -35.6995 | -38.2614
27.6523 | -35.6998 | -38.2663
27.6584 | -35.7002 | -38.2717
27.6645 | -35.7005 | -38.277
27.6706 | -35.7008 | -38.2826
27.6767 | -35.7011 | -38.2877
27.6828 | -35.7015 | -38.2927
27.6889 | -35.7019 | -38.298
27.695 | -35.7022 | -38.3032
27.7011 | -35.7025 | -38.3085
27.7072 | -35.7029 | -38.3141
27.7133 | -35.7032 | -38.3194
27.7194 | -35.7036 | -38.3242
27.7256 | -35.7039 | -38.3298
27.7317 | -35.7043 | -38.3355
27.7378 | -35.7046 | -38.3412
27.7439 | -35.7049 | -38.3466
27.75 | -35.7052 | -38.3525
27.7561 | -35.7056 | -38.3572
27.7622 | -35.7059 | -38.3623
27.7683 | -35.7062 | -38.3679
27.7744 | -35.7066 | -38.373
27.7805 | -35.7069 | -38.3785
27.7866 | -35.7072 | -38.3836
27.7927 | -35.7076 | -38.3886
27.7988 | -35.7079 | -38.3941
27.8049 | -35.7083 | -38.3993
27.811 | -35.7087 | -38.4047
27.8171 | -35.7091 | -38.41
27.8232 | -35.7095 | -38.4149
27.8293 | -35.7099 | -38.4203
27.8354 | -35.7103 | -38.4258
27.8415 | -35.7106 | -38.4315
27.8476 | -35.711 | -38.4371
27.8537 | -35.7113 | -38.4428
27.8598 | -35.7117 | -38.4477
27.866 | -35.712 | -38.4527
27.8721 | -35.7124 | -38.4578
27.8782 | -35.7127 | -38.4634
27.8843 | -35.7131 | -38.4689
27.8904 | -35.7134 | -38.4741
27.8965 | -35.7138 | -38.4791
27.9026 | -35.7141 | -38.4841
27.9087 | -35.7144 | -38.4894
27.9148 | -35.7148 | -38.4944
27.9209 | -35.7152 | -38.5
27.927 | -35.7155 | -38.5052
27.9331 | -35.7159 | -38.5102
27.9392 | -35.7163 | -38.5154
27.9453 | -35.7167 | -38.5204
27.9514 | -35.7171 | -38.5254
27.9575 | -35.7174 | -38.5304
27.9636 | -35.7177 | -38.5363
27.9697 | -35.718 | -38.5413
27.9758 | -35.7183 | -38.5465
27.9819 | -35.7187 | -38.5524
27.988 | -35.719 | -38.5572
27.9941 | -35.7193 | -38.5627
28.0002 | -35.7197 | -38.5678
28.0063 | -35.72 | -38.573
28.0125 | -35.7204 | -38.5782
28.0186 | -35.7206 | -38.584
28.0247 | -35.721 | -38.5892
28.0308 | -35.7214 | -38.5939
28.0369 | -35.7217 | -38.5986
28.043 | -35.7219 | -38.6042
28.0491 | -35.7223 | -38.6099
28.0552 | -35.7226 | -38.6148
28.0613 | -35.7228 | -38.6204
28.0674 | -35.7232 | -38.6256
28.0735 | -35.7235 | -38.6303

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.378 | -42.4283
21.4908 | -35.4701 | -35.6625
16.4948 | -34.8473 | -28.1342
12.6602 | -34.4926 | -21.7533
9.7171 | -34.2748 | -16.7611
7.4582 | -34.1269 | -12.7293
5.7244 | -34.0237 | -9.4574
4.3936 | -33.9502 | -6.8287
3.3722 | -33.9325 | -4.747
2.5883 | -33.9309 | -3.1711
1.9866 | -33.9405 | -1.9148
1.5248 | -33.9516 | -0.90647
1.1703 | -33.964 | -0.24568
1 | -33.964 | 0.081521
0.89822 | -33.9479 | 0.3563
0.68941 | -33.9306 | 0.94103
0.52919 | -33.9468 | 1.3974
0.40612 | -33.9353 | 2.0109
0.31171 | -33.9298 | 2.464
0.23928 | -33.9434 | 3.1923
0.18358 | -33.9524 | 3.6157
0.14091 | -33.9592 | 4.5043
0.10819 | -34.0039 | 5.6631
0.083017 | -34.0456 | 7.1427
0.063705 | -34.0642 | 9.0293
0.048923 | -34.1296 | 11.5291
0.037479 | -34.2401 | 14.732
0.028801 | -34.4466 | 18.7685
0.022078 | -34.7557 | 24.1262
0.017023 | -35.2472 | 30.5528
0.013018 | -36.0967 | 38.2271
0.010014 | -37.3262 | 46.1336

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003673 | -12.4564
0.007347 | -34.5517
0.01102 | -39.849
0.014693 | -42.3947
0.018366 | -45.2343
0.02204 | -47.0232
0.025713 | -48.1195
0.029386 | -48.4094
0.033059 | -50.4047
0.036733 | -52.207
0.040406 | -52.1822
0.044079 | -54.7011
0.047752 | -56.178
0.051426 | -56.6013
0.055099 | -57.0492
0.058772 | -57.8331
0.062445 | -57.6572
0.066119 | -57.8752
0.069792 | -58.7561
0.073465 | -60.2525
0.077139 | -62.3289
0.080812 | -63.2117
0.084485 | -62.8421
0.088158 | -63.178
0.091832 | -63.7957
0.095505 | -63.7915
0.099178 | -64.0005
0.10285 | -64.5028
0.10652 | -64.0106
0.1102 | -64.0769
0.11387 | -65.2443
0.11754 | -66.0563
0.12122 | -65.8799
0.12489 | -67.1258
0.12856 | -67.5826
0.13224 | -67.3324
0.13591 | -67.3019
0.13958 | -67.2838
0.14326 | -67.7005
0.14693 | -68.7962
0.1506 | -69.217
0.15428 | -68.8897
0.15795 | -69.2194
0.16162 | -69.3097
0.1653 | -69.1613
0.16897 | -69.8293
0.17264 | -70.4314
0.17632 | -70.5328
0.17999 | -71.6282
0.18366 | -70.4644
0.18734 | -70.3589
0.19101 | -70.2223
0.19468 | -69.7961
0.19836 | -70.2097
0.20203 | -71.3734
0.2057 | -71.3171
0.20938 | -71.6338
0.21305 | -72.2107
0.21672 | -72.4265
0.2204 | -72.0689
0.22407 | -71.878
0.22774 | -72.2601
0.23142 | -72.696
0.23509 | -72.731
0.23876 | -73.1786
0.24244 | -73.7813
0.24611 | -74.3116
0.24978 | -73.741
0.25345 | -73.6827
0.25713 | -74.5451
0.2608 | -74.3339
0.26448 | -74.6351
0.26815 | -73.642
0.27182 | -74.2876
0.27549 | -75.2598
0.27917 | -75.0774
0.28284 | -75.459
0.28652 | -75.5345
0.29019 | -75.1977
0.29386 | -75.8611
0.29753 | -75.655
0.30121 | -75.9176
0.30488 | -76.5247
0.30855 | -76.5604
0.31223 | -76.6782
0.3159 | -77.5037
0.31957 | -76.2699
0.32325 | -76.4688
0.32692 | -76.3934
0.33059 | -76.0346
0.33427 | -76.6324
0.33794 | -77.7587
0.34161 | -77.8156
0.34529 | -77.3968
0.34896 | -77.11
0.35263 | -77.7704
0.35631 | -78.704
0.35998 | -79.1209
0.36365 | -79.0184
0.36733 | -78.8481
0.371 | -77.6368
0.37467 | -78.3865
0.37835 | -79.1912
0.38202 | -77.9781
0.38569 | -78.4586
0.38937 | -79.1287
0.39304 | -79.8854
0.39671 | -80.0188
0.40039 | -79.5611
0.40406 | -79.5323
0.40773 | -78.3318
0.41141 | -77.9894
0.41508 | -79.1498
0.41875 | -80.0558
0.42242 | -79.568
0.4261 | -79.7565
0.42977 | -80.2812
0.43345 | -80.1669
0.43712 | -81.1661
0.44079 | -80.6925
0.44446 | -80.9106
0.44814 | -80.0475
0.45181 | -79.5748
0.45548 | -80.0562
0.45916 | -80.3526
0.46283 | -80.6546
0.4665 | -81.4859
0.47018 | -81.2764
0.47385 | -80.9036
0.47752 | -80.1613
0.4812 | -80.3511
0.48487 | -79.564
0.48854 | -79.16
0.49222 | -81.1069
0.49589 | -81.9389
0.49956 | -81.3921
0.50324 | -81.3209
0.50691 | -81.2007
0.51058 | -80.411
0.51426 | -80.7199
0.51793 | -81.5688
0.5216 | -81.8769
0.52528 | -80.8898
0.52895 | -80.7534
0.53262 | -81.0783
0.5363 | -81.8249
0.53997 | -82.3454
0.54364 | -82.3823
0.54732 | -82.5087
0.55099 | -82.4922
0.55466 | -81.9961
0.55834 | -81.5088
0.56201 | -81.5816
0.56568 | -82.1549
0.56936 | -82.4532
0.57303 | -82.4612
0.5767 | -82.1496
0.58038 | -81.6308
0.58405 | -81.4492
0.58772 | -81.649
0.5914 | -81.8207
0.59507 | -81.9655
0.59874 | -82.3098
0.60242 | -82.8013
0.60609 | -83.3805
0.60976 | -83.3815
0.61343 | -83.1032
0.61711 | -83.1368
0.62078 | -82.9704
0.62445 | -82.6361
0.62813 | -82.4917
0.6318 | -82.3583
0.63548 | -82.5196
0.63915 | -82.8837
0.64282 | -82.9766
0.64649 | -83.0641
0.65017 | -83.47
0.65384 | -83.6445
0.65751 | -83.5198
0.66119 | -83.6056
0.66486 | -84.0061
0.66853 | -84.179
0.67221 | -83.9934
0.67588 | -83.8612
0.67955 | -83.9676
0.68323 | -83.8188
0.6869 | -83.4831
0.69057 | -83.3559
0.69425 | -83.4994
0.69792 | -83.9063
0.70159 | -84.0881
0.70527 | -83.5471
0.70894 | -83.3056
0.71261 | -83.5976
0.71629 | -83.9414
0.71996 | -84.1824
0.72363 | -84.2961
0.72731 | -84.0475
0.73098 | -84.0794
0.73465 | -84.5154
0.73833 | -84.8345
0.742 | -84.8544
0.74567 | -84.6882
0.74935 | -84.329
0.75302 | -84.2334
0.75669 | -84.5322
0.76036 | -84.7634
0.76404 | -84.9603
0.76771 | -85.1778
0.77138 | -85.3019
0.77506 | -85.1464
0.77873 | -84.9124
0.78241 | -84.9902
0.78608 | -85.438
0.78975 | -85.7513
0.79343 | -85.5447
0.7971 | -85.2566
0.80077 | -85.1788
0.80444 | -84.94
0.80812 | -84.8477
0.81179 | -85.2461
0.81546 | -85.7287
0.81914 | -86.0513
0.82281 | -86.3398
0.82648 | -86.4176
0.83016 | -86.0081
0.83383 | -85.7396
0.8375 | -85.8646
0.84118 | -85.5415
0.84485 | -85.1548
0.84852 | -85.5027
0.8522 | -86.123
0.85587 | -86.2862
0.85954 | -86.3149
0.86322 | -86.0632
0.86689 | -85.9531
0.87056 | -86.1512
0.87424 | -85.832
0.87791 | -85.4886
0.88158 | -85.4878
0.88526 | -86.0171
0.88893 | -86.6673
0.8926 | -86.8951
0.89628 | -87.1034
0.89995 | -87.2078
0.90362 | -87.1871
0.9073 | -87.1923
0.91097 | -87.1586
0.91464 | -86.9181
0.91832 | -86.9879
0.92199 | -87.2137
0.92566 | -87.1261
0.92934 | -87.1502
0.93301 | -87.2866
0.93668 | -87.2356
0.94036 | -87.1844
0.94403 | -87.3237
0.9477 | -87.6181
0.95137 | -87.4407
0.95505 | -87.2721
0.95872 | -86.8206
0.9624 | -85.9609
0.96607 | -85.6634
0.96974 | -86.0441
0.97342 | -86.7033
0.97709 | -87.2092
0.98076 | -87.149
0.98443 | -86.8362
0.98811 | -86.7668
0.99178 | -87.0456
0.99545 | -87.5682
0.99913 | -87.8027
1.0028 | -87.6603
1.0065 | -87.6052
1.0101 | -87.4077
1.0138 | -87.1524
1.0175 | -87.065
1.0212 | -87.5544
1.0248 | -88.2859
1.0285 | -88.2983
1.0322 | -87.5426
1.0359 | -87.2175
1.0395 | -87.3766
1.0432 | -87.5331
1.0469 | -87.6695
1.0506 | -87.8377
1.0542 | -87.6733
1.0579 | -87.5227
1.0616 | -87.4844
1.0652 | -87.5817
1.0689 | -87.7961
1.0726 | -88.0466
1.0763 | -88.3066
1.0799 | -88.5072
1.0836 | -88.4609
1.0873 | -88.2382
1.091 | -88.1911
1.0946 | -88.4337
1.0983 | -88.705
1.102 | -88.725
1.1057 | -88.6059
1.1093 | -88.5671
1.113 | -88.6232
1.1167 | -88.6501
1.1203 | -88.6435
1.124 | -88.7195
1.1277 | -88.871
1.1314 | -88.9287
1.135 | -88.9066
1.1387 | -88.968
1.1424 | -89.0823
1.1461 | -89.1775
1.1497 | -89.3281
1.1534 | -89.3714
1.1571 | -89.0677
1.1608 | -88.6164
1.1644 | -88.2938
1.1681 | -88.0831
1.1718 | -87.9009
1.1754 | -87.8234
1.1791 | -88.0007
1.1828 | -88.3772
1.1865 | -88.7193
1.1901 | -88.958
1.1938 | -89.1082
1.1975 | -89.0356
1.2012 | -88.8131
1.2048 | -88.7164
1.2085 | -88.7653
1.2122 | -88.8074
1.2159 | -88.9489
1.2195 | -89.3667
1.2232 | -89.9229
1.2269 | -90.297
1.2305 | -90.3665
1.2342 | -90.1415
1.2379 | -89.7437
1.2416 | -89.3875
1.2452 | -89.1959
1.2489 | -89.2286
1.2526 | -89.3666
1.2563 | -89.2921
1.2599 | -88.9721
1.2636 | -88.7299
1.2673 | -88.7326
1.2709 | -88.857
1.2746 | -89.1197
1.2783 | -89.7165
1.282 | -90.4632
1.2856 | -90.7415
1.2893 | -90.3925
1.293 | -89.9081
1.2967 | -89.6488
1.3003 | -89.6632
1.304 | -89.8094
1.3077 | -89.8762
1.3114 | -89.7287
1.315 | -89.5021
1.3187 | -89.4872
1.3224 | -89.6758
1.326 | -89.9052
1.3297 | -90.1367
1.3334 | -90.3043
1.3371 | -90.337
1.3407 | -90.211
1.3444 | -89.9839
1.3481 | -89.7977
1.3518 | -89.7424
1.3554 | -89.8253
1.3591 | -89.9669
1.3628 | -90.127
1.3665 | -90.3436
1.3701 | -90.534
1.3738 | -90.5948
1.3775 | -90.4693
1.3811 | -90.2819
1.3848 | -90.2519
1.3885 | -90.3862
1.3922 | -90.5777
1.3958 | -90.7249
1.3995 | -90.6516
1.4032 | -90.3778
1.4069 | -90.1768
1.4105 | -90.2816
1.4142 | -90.6415
1.4179 | -90.9387
1.4216 | -90.8745
1.4252 | -90.5725
1.4289 | -90.4593
1.4326 | -90.6888
1.4362 | -90.9759
1.4399 | -91.0139
1.4436 | -90.862
1.4473 | -90.7641
1.4509 | -90.7806
1.4546 | -90.8043
1.4583 | -90.8157
1.462 | -90.7757
1.4656 | -90.6803
1.4693 | -90.7205
1.473 | -91.0064
1.4767 | -91.3748
1.4803 | -91.6209
1.484 | -91.7982
1.4877 | -91.9931
1.4913 | -92.0232
1.495 | -91.8038
1.4987 | -91.5696
1.5024 | -91.4368
1.506 | -91.2608
1.5097 | -91.0874
1.5134 | -91.0896
1.5171 | -91.1435
1.5207 | -91.0612
1.5244 | -90.9736
1.5281 | -91.0795
1.5318 | -91.307
1.5354 | -91.3809
1.5391 | -91.3245
1.5428 | -91.3699
1.5464 | -91.5848
1.5501 | -91.9403
1.5538 | -92.2757
1.5575 | -92.3429
1.5611 | -92.1177
1.5648 | -91.7767
1.5685 | -91.5122
1.5722 | -91.3221
1.5758 | -91.0486
1.5795 | -90.7373
1.5832 | -90.6979
1.5869 | -91.0648
1.5905 | -91.2884
1.5942 | -91.3803
1.5979 | -91.3139
1.6015 | -91.2383
1.6052 | -91.2787
1.6089 | -91.4301
1.6126 | -91.5879
1.6162 | -91.7193
1.6199 | -91.8883
1.6236 | -92.1402
1.6273 | -92.3812
1.6309 | -92.4286
1.6346 | -92.205
1.6383 | -91.8494
1.6419 | -91.5511
1.6456 | -91.4022
1.6493 | -91.4222
1.653 | -91.58
1.6566 | -91.8285
1.6603 | -92.1569
1.664 | -92.509
1.6677 | -92.7455
1.6713 | -92.7845
1.675 | -92.7177
1.6787 | -92.662
1.6824 | -92.6033
1.686 | -92.4924
1.6897 | -92.3334
1.6934 | -92.1528
1.697 | -91.9704
1.7007 | -91.8255
1.7044 | -91.7608
1.7081 | -91.7828
1.7117 | -91.8767
1.7154 | -92.0243
1.7191 | -92.2012
1.7228 | -92.371
1.7264 | -92.4774
1.7301 | -92.4383
1.7338 | -92.2489
1.7375 | -92.0131
1.7411 | -91.8231
1.7448 | -91.7076
1.7485 | -91.6505
1.7521 | -91.6242
1.7558 | -91.5934
1.7595 | -91.5982
1.7632 | -91.7239
1.7668 | -91.9552
1.7705 | -92.2079
1.7742 | -92.3999
1.7779 | -92.5101
1.7815 | -92.5588
1.7852 | -92.5726
1.7889 | -92.583
1.7926 | -92.5795
1.7962 | -92.5666
1.7999 | -92.6251
1.8036 | -92.7808
1.8072 | -92.9395
1.8109 | -93.0101
1.8146 | -93.023
1.8183 | -93.0269
1.8219 | -92.9619
1.8256 | -92.7939
1.8293 | -92.6147
1.833 | -92.4935
1.8366 | -92.419
1.8403 | -92.3535
1.844 | -92.3111
1.8477 | -92.3485
1.8513 | -92.4642
1.855 | -92.6035
1.8587 | -92.7371
1.8623 | -92.8899
1.866 | -93.0884
1.8697 | -93.283
1.8734 | -93.3926
1.877 | -93.3922
1.8807 | -93.3179
1.8844 | -93.2108
1.8881 | -93.0399
1.8917 | -92.8119
1.8954 | -92.6351
1.8991 | -92.629
1.9027 | -92.7848
1.9064 | -92.9573
1.9101 | -92.9858
1.9138 | -92.7742
1.9174 | -92.3643
1.9211 | -91.9742
1.9248 | -91.7939
1.9285 | -91.8708
1.9321 | -92.1394
1.9358 | -92.5634
1.9395 | -93.1055
1.9432 | -93.6345
1.9468 | -93.9488
1.9505 | -93.9115
1.9542 | -93.6443
1.9578 | -93.4022
1.9615 | -93.2931
1.9652 | -93.2517
1.9689 | -93.1787
1.9725 | -93.1034
1.9762 | -93.1172
1.9799 | -93.2505
1.9836 | -93.4452
1.9872 | -93.6342
1.9909 | -93.8125
1.9946 | -93.9564
1.9983 | -94.0029
2.0019 | -93.9192
2.0056 | -93.7577
2.0093 | -93.5961
2.0129 | -93.4587
2.0166 | -93.3316
2.0203 | -93.2032
2.024 | -93.1094
2.0276 | -93.1074
2.0313 | -93.2061
2.035 | -93.3757
2.0387 | -93.5704
2.0423 | -93.7465
2.046 | -93.8558
2.0497 | -93.8376
2.0534 | -93.7338
2.057 | -93.6592
2.0607 | -93.6712
2.0644 | -93.7547
2.068 | -93.8468
2.0717 | -93.9008
2.0754 | -93.8958
2.0791 | -93.8523
2.0827 | -93.7987
2.0864 | -93.7389
2.0901 | -93.6543
2.0938 | -93.552
2.0974 | -93.4778
2.1011 | -93.4974
2.1048 | -93.6241
2.1085 | -93.8482
2.1121 | -94.1309
2.1158 | -94.3757
2.1195 | -94.3502
2.1231 | -94.3533
2.1268 | -94.3457
2.1305 | -94.3817
2.1342 | -94.4478
2.1378 | -94.4625
2.1415 | -94.3466
2.1452 | -94.1184
2.1489 | -93.8723
2.1525 | -93.7197
2.1562 | -93.7176
2.1599 | -93.8509
2.1636 | -94.0344
2.1672 | -94.1922
2.1709 | -94.301
2.1746 | -94.3412
2.1782 | -94.3134
2.1819 | -94.2533
2.1856 | -94.2125
2.1893 | -94.2345
2.1929 | -94.3291
2.1966 | -94.4464
2.2003 | -94.529
2.204 | -94.5332
2.2076 | -94.4541
2.2113 | -94.3397
2.215 | -94.2548
2.2187 | -94.2372
2.2223 | -94.2809
2.226 | -94.3562
2.2297 | -94.4311
2.2333 | -94.4529
2.237 | -94.4316
2.2407 | -94.4088
2.2444 | -94.4319
2.248 | -94.5282
2.2517 | -94.6923
2.2554 | -94.8677
2.2591 | -94.9642
2.2627 | -94.9013
2.2664 | -94.6353
2.2701 | -94.2689
2.2738 | -93.9572
2.2774 | -93.7917
2.2811 | -93.7773
2.2848 | -93.8589
2.2884 | -93.9904
2.2921 | -94.1596
2.2958 | -94.3698
2.2995 | -94.5832
2.3031 | -94.7539
2.3068 | -94.8551
2.3105 | -94.8743
2.3142 | -94.8425
2.3178 | -94.769
2.3215 | -94.6645
2.3252 | -94.5335
2.3288 | -94.383
2.3325 | -94.2534
2.3362 | -94.1733
2.3399 | -94.1636
2.3435 | -94.2207
2.3472 | -94.3505
2.3509 | -94.5313
2.3546 | -94.7429
2.3582 | -94.9318
2.3619 | -95.0333
2.3656 | -95.0667
2.3693 | -95.0881
2.3729 | -95.1446
2.3766 | -95.252
2.3803 | -95.3773
2.3839 | -95.4638
2.3876 | -95.4628
2.3913 | -95.3986
2.395 | -95.2943
2.3986 | -95.1982
2.4023 | -95.149
2.406 | -95.1639
2.4097 | -95.2742
2.4133 | -95.4679
2.417 | -95.6888
2.4207 | -95.8521
2.4244 | -95.9217
2.428 | -95.8842
2.4317 | -95.7732
2.4354 | -95.6296
2.439 | -95.4908
2.4427 | -95.3491
2.4464 | -95.216
2.4501 | -95.0915
2.4537 | -94.9803
2.4574 | -94.881
2.4611 | -94.7811
2.4648 | -94.6907
2.4684 | -94.6481
2.4721 | -94.6947
2.4758 | -94.851
2.4795 | -95.0579
2.4831 | -95.2342
2.4868 | -95.3239
2.4905 | -95.337
2.4941 | -95.2958
2.4978 | -95.2352
2.5015 | -95.2088
2.5052 | -95.2482
2.5088 | -95.3706
2.5125 | -95.5418
2.5162 | -95.6877
2.5199 | -95.721
2.5235 | -95.5918
2.5272 | -95.3365
2.5309 | -95.0463
2.5346 | -94.8097
2.5382 | -94.6668
2.5419 | -94.6345
2.5456 | -94.7229
2.5492 | -94.8945
2.5529 | -95.1029
2.5566 | -95.3398
2.5603 | -95.6076
2.5639 | -95.861
2.5676 | -96.0302
2.5713 | -96.1121
2.575 | -96.1208
2.5786 | -96.0877
2.5823 | -96.0225
2.586 | -95.9323
2.5897 | -95.8616
2.5933 | -95.8116
2.597 | -95.7502
2.6007 | -95.6633
2.6043 | -95.5703
2.608 | -95.4754
2.6117 | -95.354
2.6154 | -95.2121
2.619 | -95.0796
2.6227 | -95.0047
2.6264 | -95.0128
2.6301 | -95.0729
2.6337 | -95.1625
2.6374 | -95.2868
2.6411 | -95.4599
2.6448 | -95.6719
2.6484 | -95.8288
2.6521 | -95.9418
2.6558 | -95.9672
2.6594 | -95.8924
2.6631 | -95.7295
2.6668 | -95.53
2.6705 | -95.3527
2.6741 | -95.2375
2.6778 | -95.185
2.6815 | -95.1732
2.6852 | -95.1827
2.6888 | -95.1905
2.6925 | -95.1746
2.6962 | -95.1419
2.6998 | -95.1152
2.7035 | -95.1352
2.7072 | -95.2283
2.7109 | -95.4006
2.7145 | -95.6171
2.7182 | -95.8125
2.7219 | -95.9348
2.7256 | -95.9561
2.7292 | -95.8777
2.7329 | -95.7384
2.7366 | -95.5827
2.7403 | -95.4435
2.7439 | -95.3271
2.7476 | -95.2475
2.7513 | -95.2294
2.7549 | -95.2715
2.7586 | -95.3681
2.7623 | -95.4928
2.766 | -95.6098
2.7696 | -95.6821
2.7733 | -95.7156
2.777 | -95.7453
2.7807 | -95.8018
2.7843 | -95.9043
2.788 | -96.0346
2.7917 | -96.1714
2.7954 | -96.2784
2.799 | -96.3314
2.8027 | -96.3476
2.8064 | -96.3247
2.81 | -96.2921
2.8137 | -96.2643
2.8174 | -96.2572
2.8211 | -96.2658
2.8247 | -96.2812
2.8284 | -96.2712
2.8321 | -96.2443
2.8358 | -96.1986
2.8394 | -96.1327
2.8431 | -96.08
2.8468 | -96.0652
2.8505 | -96.1058
2.8541 | -96.2143
2.8578 | -96.3574
2.8615 | -96.4805
2.8651 | -96.5049
2.8688 | -96.3928
2.8725 | -96.1639
2.8762 | -95.8937
2.8798 | -95.6522
2.8835 | -95.5018
2.8872 | -95.4751
2.8909 | -95.554
2.8945 | -95.7121
2.8982 | -95.8919
2.9019 | -96.0452
2.9056 | -96.1324
2.9092 | -96.1452
2.9129 | -96.1047
2.9166 | -96.0366
2.9202 | -95.9702
2.9239 | -95.9324
2.9276 | -95.9353
2.9313 | -95.9813
2.9349 | -96.065
2.9386 | -96.1786
2.9423 | -96.2942
2.946 | -96.416
2.9496 | -96.5434
2.9533 | -96.6752
2.957 | -96.806
2.9606 | -96.9123
2.9643 | -96.9679
2.968 | -96.9427
2.9717 | -96.8224
2.9753 | -96.6212
2.979 | -96.3849
2.9827 | -96.1623
2.9864 | -96.0034
2.99 | -95.9239
2.9937 | -95.9089
2.9974 | -95.923
3.0011 | -95.9447
3.0047 | -95.9615
3.0084 | -95.9852
3.0121 | -96.0399
3.0157 | -96.1447
3.0194 | -96.2814
3.0231 | -96.4229
3.0268 | -96.539
3.0304 | -96.5814
3.0341 | -96.5383
3.0378 | -96.4257
3.0415 | -96.2755
3.0451 | -96.1238
3.0488 | -96.0009
3.0525 | -95.9231
3.0562 | -95.906
3.0598 | -95.9576
3.0635 | -96.0557
3.0672 | -96.172
3.0708 | -96.2602
3.0745 | -96.2901
3.0782 | -96.261
3.0819 | -96.1786
3.0855 | -96.0827
3.0892 | -96.0312
3.0929 | -96.0829
3.0966 | -96.2516
3.1002 | -96.5111
3.1039 | -96.7843
3.1076 | -96.9967
3.1113 | -97.1479
3.1149 | -97.254
3.1186 | -97.348
3.1223 | -97.4195
3.1259 | -97.444
3.1296 | -97.4052
3.1333 | -97.3161
3.137 | -97.2026
3.1406 | -97.0877
3.1443 | -96.9886
3.148 | -96.9134
3.1517 | -96.8669
3.1553 | -96.8415
3.159 | -96.8298
3.1627 | -96.8535
3.1664 | -96.914
3.17 | -97.0039
3.1737 | -97.0941
3.1774 | -97.0851
3.181 | -97.0903
3.1847 | -97.0798
3.1884 | -97.081
3.1921 | -97.1109
3.1957 | -97.1669
3.1994 | -97.2167
3.2031 | -97.2506
3.2068 | -97.2458
3.2104 | -97.1958
3.2141 | -97.1006
3.2178 | -96.9803
3.2215 | -96.8777
3.2251 | -96.8202
3.2288 | -96.825
3.2325 | -96.8936
3.2361 | -97.0032
3.2398 | -97.1253
3.2435 | -97.247
3.2472 | -97.3433
3.2508 | -97.4038
3.2545 | -97.4193
3.2582 | -97.3692
3.2619 | -97.2684
3.2655 | -97.1333
3.2692 | -96.9858
3.2729 | -96.8472
3.2766 | -96.7417
3.2802 | -96.6908
3.2839 | -96.706
3.2876 | -96.7837
3.2912 | -96.9086
3.2949 | -97.0562
3.2986 | -97.1944
3.3023 | -97.3038
3.3059 | -97.3674
3.3096 | -97.384
3.3133 | -97.362
3.317 | -97.2983
3.3206 | -97.2144
3.3243 | -97.1374
3.328 | -97.0782
3.3316 | -97.0427
3.3353 | -97.0416
3.339 | -97.0653
3.3427 | -97.1001
3.3463 | -97.1374
3.35 | -97.1649
3.3537 | -97.1876
3.3574 | -97.2067
3.361 | -97.2302
3.3647 | -97.2626
3.3684 | -97.3167
3.3721 | -97.3916
3.3757 | -97.467
3.3794 | -97.5183
3.3831 | -97.5399
3.3867 | -97.5222
3.3904 | -97.4711
3.3941 | -97.4043
3.3978 | -97.3405
3.4014 | -97.2969
3.4051 | -97.2929
3.4088 | -97.3348
3.4125 | -97.4261
3.4161 | -97.5453
3.4198 | -97.6671
3.4235 | -97.763
3.4272 | -97.8314
3.4308 | -97.8877
3.4345 | -97.9516
3.4382 | -98.0306
3.4418 | -98.126
3.4455 | -98.2299
3.4492 | -98.3122
3.4529 | -98.341
3.4565 | -98.3018
3.4602 | -98.1926
3.4639 | -98.0184
3.4676 | -97.8096
3.4712 | -97.6185
3.4749 | -97.4809
3.4786 | -97.4149
3.4823 | -97.4309
3.4859 | -97.5096
3.4896 | -97.6205
3.4933 | -97.7233
3.4969 | -97.7825
3.5006 | -97.7882
3.5043 | -97.7467
3.508 | -97.6583
3.5116 | -97.5252
3.5153 | -97.3712
3.519 | -97.2195
3.5227 | -97.0874
3.5263 | -96.9962
3.53 | -96.9717
3.5337 | -97.0109
3.5374 | -97.0974
3.541 | -97.2108
3.5447 | -97.3401
3.5484 | -97.4699
3.552 | -97.5927
3.5557 | -97.6949
3.5594 | -97.7846
3.5631 | -97.8637
3.5667 | -97.9293
3.5704 | -97.967
3.5741 | -97.9935
3.5778 | -97.9984
3.5814 | -97.9716
3.5851 | -97.9156
3.5888 | -97.8373
3.5925 | -97.7597
3.5961 | -97.7019
3.5998 | -97.6747
3.6035 | -97.6874
3.6071 | -97.7334
3.6108 | -97.8043
3.6145 | -97.8879
3.6182 | -97.9811
3.6218 | -98.0913
3.6255 | -98.2063
3.6292 | -98.3011
3.6329 | -98.3818
3.6365 | -98.4346
3.6402 | -98.4497
3.6439 | -98.4315
3.6475 | -98.3608
3.6512 | -98.2679
3.6549 | -98.1676
3.6586 | -98.0759
3.6622 | -98.0107
3.6659 | -97.9778
3.6696 | -97.9725
3.6733 | -97.9931
3.6769 | -98.0236
3.6806 | -98.0489
3.6843 | -98.0506
3.688 | -98.0315
3.6916 | -98.0171
3.6953 | -98.0349
3.699 | -98.0994
3.7026 | -98.1999
3.7063 | -98.2724
3.71 | -98.323
3.7137 | -98.347
3.7173 | -98.3294
3.721 | -98.2738
3.7247 | -98.1945
3.7284 | -98.1011
3.732 | -98.009
3.7357 | -97.9299
3.7394 | -97.8685
3.7431 | -97.8318
3.7467 | -97.8086
3.7504 | -97.7993
3.7541 | -97.7952
3.7577 | -97.7853
3.7614 | -97.769
3.7651 | -97.7434
3.7688 | -97.7001
3.7724 | -97.6557
3.7761 | -97.6189
3.7798 | -97.6093
3.7835 | -97.6264
3.7871 | -97.6744
3.7908 | -97.7519
3.7945 | -97.8493
3.7982 | -97.9677
3.8018 | -98.1061
3.8055 | -98.2525
3.8092 | -98.4069
3.8128 | -98.5533
3.8165 | -98.688
3.8202 | -98.8134
3.8239 | -98.924
3.8275 | -99.0073
3.8312 | -99.074
3.8349 | -99.125
3.8386 | -99.1649
3.8422 | -99.1989
3.8459 | -99.2292
3.8496 | -99.244
3.8533 | -99.2511
3.8569 | -99.2415
3.8606 | -99.2059
3.8643 | -99.1534
3.8679 | -99.0953
3.8716 | -99.0454
3.8753 | -99.008
3.879 | -98.9791
3.8826 | -98.952
3.8863 | -98.9167
3.89 | -98.8752
3.8937 | -98.8272
3.8973 | -98.7713
3.901 | -98.7252
3.9047 | -98.694
3.9084 | -98.6857
3.912 | -98.6978
3.9157 | -98.7241
3.9194 | -98.7647
3.923 | -98.8121
3.9267 | -98.8548
3.9304 | -98.8869
3.9341 | -98.9035
3.9377 | -98.911
3.9414 | -98.9192
3.9451 | -98.933
3.9488 | -98.9593
3.9524 | -98.9879
3.9561 | -99.0139
3.9598 | -99.0477
3.9635 | -99.0711
3.9671 | -99.0773
3.9708 | -99.069
3.9745 | -99.0506
3.9781 | -99.0278
3.9818 | -99.0006
3.9855 | -98.9651
3.9892 | -98.9293
3.9928 | -98.8984
3.9965 | -98.8769
4.0002 | -98.8653
4.0039 | -98.852
4.0075 | -98.8485
4.0112 | -98.8611
4.0149 | -98.8941
4.0186 | -98.931
4.0222 | -98.9626
4.0259 | -98.9933
4.0296 | -99.0324
4.0332 | -99.075
4.0369 | -99.1233
4.0406 | -99.173
4.0443 | -99.2284
4.0479 | -99.298
4.0516 | -99.3633
4.0553 | -99.4228
4.059 | -99.466
4.0626 | -99.4981
4.0663 | -99.5148
4.07 | -99.5067
4.0736 | -99.4878
4.0773 | -99.4656
4.081 | -99.4452
4.0847 | -99.4289
4.0883 | -99.4215
4.092 | -99.4112
4.0957 | -99.3915
4.0994 | -99.354
4.103 | -99.2969
4.1067 | -99.222
4.1104 | -99.1267
4.1141 | -99.0271
4.1177 | -98.9247
4.1214 | -98.8416
4.1251 | -98.7804
4.1287 | -98.7672
4.1324 | -98.8012
4.1361 | -98.8698
4.1398 | -98.9774
4.1434 | -99.1094
4.1471 | -99.2359
4.1508 | -99.3377
4.1545 | -99.4096
4.1581 | -99.4471
4.1618 | -99.4394
4.1655 | -99.3866
4.1692 | -99.3087
4.1728 | -99.2315
4.1765 | -99.1691
4.1802 | -99.1476
4.1838 | -99.1705
4.1875 | -99.2401
4.1912 | -99.3284
4.1949 | -99.4296
4.1985 | -99.5268
4.2022 | -99.5911
4.2059 | -99.6262
4.2096 | -99.642
4.2132 | -99.635
4.2169 | -99.6161
4.2206 | -99.605
4.2243 | -99.5999
4.2279 | -99.6011
4.2316 | -99.6115
4.2353 | -99.6277
4.2389 | -99.654
4.2426 | -99.6711
4.2463 | -99.6797
4.25 | -99.674
4.2536 | -99.6535
4.2573 | -99.6112
4.261 | -99.5425
4.2647 | -99.4558
4.2683 | -99.3695
4.272 | -99.2886
4.2757 | -99.2171
4.2794 | -99.1722
4.283 | -99.1549
4.2867 | -99.1576
4.2904 | -99.1936
4.294 | -99.2539
4.2977 | -99.3305
4.3014 | -99.4321
4.3051 | -99.5366
4.3087 | -99.6375
4.3124 | -99.7404
4.3161 | -99.8375
4.3198 | -99.9302
4.3234 | -100.0176
4.3271 | -100.1096
4.3308 | -100.2049
4.3345 | -100.295
4.3381 | -100.3732
4.3418 | -100.4427
4.3455 | -100.5047
4.3491 | -100.5576
4.3528 | -100.6033
4.3565 | -100.6443
4.3602 | -100.6589
4.3638 | -100.6504
4.3675 | -100.6002
4.3712 | -100.5079
4.3749 | -100.3805
4.3785 | -100.2227
4.3822 | -100.0406
4.3859 | -99.8578
4.3895 | -99.6877
4.3932 | -99.5469
4.3969 | -99.4499
4.4006 | -99.3998
4.4042 | -99.3996
4.4079 | -99.4374
4.4116 | -99.5056
4.4153 | -99.5914
4.4189 | -99.6774
4.4226 | -99.7474
4.4263 | -99.7908
4.43 | -99.8101
4.4336 | -99.8034
4.4373 | -99.7921
4.441 | -99.7755
4.4446 | -99.7678
4.4483 | -99.771
4.452 | -99.7815
4.4557 | -99.7969
4.4593 | -99.8183
4.463 | -99.8362
4.4667 | -99.8424
4.4704 | -99.8378
4.474 | -99.8196
4.4777 | -99.798
4.4814 | -99.7762
4.4851 | -99.7552
4.4887 | -99.7396
4.4924 | -99.7397
4.4961 | -99.7542
4.4997 | -99.7782
4.5034 | -99.8121
4.5071 | -99.8557
4.5108 | -99.9099
4.5144 | -99.975
4.5181 | -100.0544
4.5218 | -100.1491
4.5255 | -100.2443
4.5291 | -100.3219
4.5328 | -100.3776
4.5365 | -100.4003
4.5402 | -100.393
4.5438 | -100.3621
4.5475 | -100.3042
4.5512 | -100.2344
4.5548 | -100.1736
4.5585 | -100.1253
4.5622 | -100.0809
4.5659 | -100.0442
4.5695 | -100.0196
4.5732 | -99.9981
4.5769 | -99.9745
4.5806 | -99.9439
4.5842 | -99.9045
4.5879 | -99.8507
4.5916 | -99.7981
4.5953 | -99.7571
4.5989 | -99.7258
4.6026 | -99.7083
4.6063 | -99.7085
4.6099 | -99.7279
4.6136 | -99.7601
4.6173 | -99.8112
4.621 | -99.8738
4.6246 | -99.9334
4.6283 | -99.9862
4.632 | -100.0388
4.6357 | -100.0782
4.6393 | -100.0961
4.643 | -100.1004
4.6467 | -100.0976
4.6504 | -100.0933
4.654 | -100.0813
4.6577 | -100.0749
4.6614 | -100.0635
4.665 | -100.049
4.6687 | -100.0535
4.6724 | -100.0686
4.6761 | -100.0879
4.6797 | -100.099
4.6834 | -100.1015
4.6871 | -100.0831
4.6908 | -100.0413
4.6944 | -99.976
4.6981 | -99.8938
4.7018 | -99.7918
4.7054 | -99.6848
4.7091 | -99.5931
4.7128 | -99.518
4.7165 | -99.4622
4.7201 | -99.433
4.7238 | -99.4287
4.7275 | -99.449
4.7312 | -99.4998
4.7348 | -99.5674
4.7385 | -99.6356
4.7422 | -99.7105
4.7459 | -99.7755
4.7495 | -99.8279
4.7532 | -99.8768
4.7569 | -99.9151
4.7605 | -99.9537
4.7642 | -100.0066
4.7679 | -100.0885
4.7716 | -100.1569
4.7752 | -100.2257
4.7789 | -100.289
4.7826 | -100.3359
4.7863 | -100.3638
4.7899 | -100.3717
4.7936 | -100.3641
4.7973 | -100.3493
4.801 | -100.3237
4.8046 | -100.292
4.8083 | -100.2739
4.812 | -100.2644
4.8156 | -100.2535
4.8193 | -100.2584
4.823 | -100.2707
4.8267 | -100.2817
4.8303 | -100.2904
4.834 | -100.2922
4.8377 | -100.2849
4.8414 | -100.2747
4.845 | -100.2589
4.8487 | -100.2361
4.8524 | -100.2088
4.8561 | -100.1742
4.8597 | -100.1316
4.8634 | -100.0865
4.8671 | -100.0417
4.8707 | -99.9992
4.8744 | -99.9608
4.8781 | -99.9297
4.8818 | -99.9099
4.8854 | -99.8929
4.8891 | -99.8846
4.8928 | -99.8868
4.8965 | -99.9001
4.9001 | -99.9322
4.9038 | -99.9822
4.9075 | -100.059
4.9112 | -100.1491
4.9148 | -100.2536
4.9185 | -100.3587
4.9222 | -100.4581
4.9258 | -100.5458
4.9295 | -100.6126
4.9332 | -100.6567
4.9369 | -100.6682
4.9405 | -100.6485
4.9442 | -100.6042
4.9479 | -100.54
4.9516 | -100.4696
4.9552 | -100.4077
4.9589 | -100.3555
4.9626 | -100.3238
4.9663 | -100.3247
4.9699 | -100.3582
4.9736 | -100.4119
4.9773 | -100.4873
4.9809 | -100.5762
4.9846 | -100.6648
4.9883 | -100.7434
4.992 | -100.7987
4.9956 | -100.8328
4.9993 | -100.8426
5.003 | -100.8311
5.0067 | -100.8001
5.0103 | -100.7509
5.014 | -100.698
5.0177 | -100.6446
5.0214 | -100.596
5.025 | -100.5576
5.0287 | -100.5278
5.0324 | -100.4998
5.036 | -100.4785
5.0397 | -100.4573
5.0434 | -100.4315
5.0471 | -100.3975
5.0507 | -100.3567
5.0544 | -100.3149
5.0581 | -100.2809
5.0618 | -100.2562
5.0654 | -100.2457
5.0691 | -100.2411
5.0728 | -100.2442
5.0765 | -100.2637
5.0801 | -100.2962
5.0838 | -100.3412
5.0875 | -100.4
5.0911 | -100.471
5.0948 | -100.5476
5.0985 | -100.6305
5.1022 | -100.7227
5.1058 | -100.8192
5.1095 | -100.9161
5.1132 | -101.0043
5.1169 | -101.0761
5.1205 | -101.1275
5.1242 | -101.1639
5.1279 | -101.1811
5.1315 | -101.1764
5.1352 | -101.1515
5.1389 | -101.1128
5.1426 | -101.0693
5.1462 | -101.0323
5.1499 | -100.9979
5.1536 | -100.9696
5.1573 | -100.9515
5.1609 | -100.9426
5.1646 | -100.9461
5.1683 | -100.9609
5.172 | -100.987
5.1756 | -101.0171
5.1793 | -101.0603
5.183 | -101.1151
5.1866 | -101.1818
5.1903 | -101.2619
5.194 | -101.3476
5.1977 | -101.4326
5.2013 | -101.5078
5.205 | -101.5675
5.2087 | -101.6016
5.2124 | -101.6147
5.216 | -101.5956
5.2197 | -101.5332
5.2234 | -101.4399
5.2271 | -101.3189
5.2307 | -101.172
5.2344 | -101.03
5.2381 | -100.8914
5.2417 | -100.7596
5.2454 | -100.6505
5.2491 | -100.5615
5.2528 | -100.494
5.2564 | -100.4475
5.2601 | -100.4253
5.2638 | -100.4132
5.2675 | -100.4038
5.2711 | -100.4068
5.2748 | -100.4246
5.2785 | -100.452
5.2822 | -100.4837
5.2858 | -100.5175
5.2895 | -100.5557
5.2932 | -100.5989
5.2968 | -100.6488
5.3005 | -100.6964
5.3042 | -100.742
5.3079 | -100.7897
5.3115 | -100.8396
5.3152 | -100.8786
5.3189 | -100.9208
5.3226 | -100.9601
5.3262 | -100.9932
5.3299 | -101.0274
5.3336 | -101.0567
5.3373 | -101.0819
5.3409 | -101.1081
5.3446 | -101.1366
5.3483 | -101.169
5.3519 | -101.193
5.3556 | -101.2116
5.3593 | -101.2304
5.363 | -101.2509
5.3666 | -101.2714
5.3703 | -101.2889
5.374 | -101.3088
5.3777 | -101.325
5.3813 | -101.3444
5.385 | -101.3696
5.3887 | -101.3943
5.3924 | -101.4249
5.396 | -101.4498
5.3997 | -101.4753
5.4034 | -101.4991
5.407 | -101.5121
5.4107 | -101.5099
5.4144 | -101.4974
5.4181 | -101.4775
5.4217 | -101.4488
5.4254 | -101.4179
5.4291 | -101.3921
5.4328 | -101.3721
5.4364 | -101.3525
5.4401 | -101.3347
5.4438 | -101.3211
5.4474 | -101.3012
5.4511 | -101.2787
5.4548 | -101.2512
5.4585 | -101.212
5.4621 | -101.1748
5.4658 | -101.1304
5.4695 | -101.0859
5.4732 | -101.0459
5.4768 | -101.0136
5.4805 | -100.9964
5.4842 | -100.9892
5.4879 | -100.9969
5.4915 | -101.0265
5.4952 | -101.0689
5.4989 | -101.1128
5.5025 | -101.1597
5.5062 | -101.2062
5.5099 | -101.2444
5.5136 | -101.2762
5.5172 | -101.3018
5.5209 | -101.3274
5.5246 | -101.3442
5.5283 | -101.3625
5.5319 | -101.3916
5.5356 | -101.4169
5.5393 | -101.4415
5.543 | -101.4694
5.5466 | -101.4933
5.5503 | -101.5193
5.554 | -101.5478
5.5576 | -101.5707
5.5613 | -101.5949
5.565 | -101.6286
5.5687 | -101.6705
5.5723 | -101.7106
5.576 | -101.7467
5.5797 | -101.7818
5.5834 | -101.8124
5.587 | -101.8238
5.5907 | -101.8159
5.5944 | -101.7997
5.5981 | -101.769
5.6017 | -101.7244
5.6054 | -101.6709
5.6091 | -101.6191
5.6127 | -101.5678
5.6164 | -101.5139
5.6201 | -101.4668
5.6238 | -101.4261
5.6274 | -101.3854
5.6311 | -101.3504
5.6348 | -101.327
5.6385 | -101.3164
5.6421 | -101.3245
5.6458 | -101.3455
5.6495 | -101.3789
5.6532 | -101.4287
5.6568 | -101.4892
5.6605 | -101.5538
5.6642 | -101.6162
5.6678 | -101.6736
5.6715 | -101.7219
5.6752 | -101.7573
5.6789 | -101.7822
5.6825 | -101.7965
5.6862 | -101.7997
5.6899 | -101.7904
5.6936 | -101.7685
5.6972 | -101.7266
5.7009 | -101.6733
5.7046 | -101.6164
5.7083 | -101.5514
5.7119 | -101.4837
5.7156 | -101.4107
5.7193 | -101.3401
5.7229 | -101.2768
5.7266 | -101.2287
5.7303 | -101.1887
5.734 | -101.1613
5.7376 | -101.1507
5.7413 | -101.1673
5.745 | -101.206
5.7487 | -101.2542
5.7523 | -101.3078
5.756 | -101.3715
5.7597 | -101.4334
5.7633 | -101.4805
5.767 | -101.5141
5.7707 | -101.5215
5.7744 | -101.5042
5.778 | -101.4674
5.7817 | -101.4275
5.7854 | -101.3853
5.7891 | -101.3463
5.7927 | -101.3217
5.7964 | -101.3107
5.8001 | -101.3187
5.8038 | -101.3481
5.8074 | -101.3969
5.8111 | -101.4588
5.8148 | -101.535
5.8184 | -101.6126
5.8221 | -101.6816
5.8258 | -101.6931
5.8295 | -101.7216
5.8331 | -101.7367
5.8368 | -101.744
5.8405 | -101.7359
5.8442 | -101.7049
5.8478 | -101.6614
5.8515 | -101.6074
5.8552 | -101.5537
5.8589 | -101.5057
5.8625 | -101.454
5.8662 | -101.4016
5.8699 | -101.3496
5.8735 | -101.3052
5.8772 | -101.266
5.8809 | -101.2301
5.8846 | -101.2022
5.8882 | -101.1747
5.8919 | -101.1517
5.8956 | -101.1315
5.8993 | -101.1153
5.9029 | -101.1017
5.9066 | -101.0922
5.9103 | -101.0861
5.914 | -101.0815
5.9176 | -101.0821
5.9213 | -101.0815
5.925 | -101.0805
5.9286 | -101.0856
5.9323 | -101.0995
5.936 | -101.1143
5.9397 | -101.1291
5.9433 | -101.1476
5.947 | -101.166
5.9507 | -101.1886
5.9544 | -101.2216
5.958 | -101.2621
5.9617 | -101.3126
5.9654 | -101.3689
5.9691 | -101.4328
5.9727 | -101.509
5.9764 | -101.5912
5.9801 | -101.6766
5.9837 | -101.7596
5.9874 | -101.8468
5.9911 | -101.935
5.9948 | -102.0141
5.9984 | -102.087
6.0021 | -102.1485
6.0058 | -102.1998
6.0095 | -102.2358
6.0131 | -102.2602
6.0168 | -102.2789
6.0205 | -102.2845
6.0242 | -102.2781
6.0278 | -102.2718
6.0315 | -102.2664
6.0352 | -102.2537
6.0388 | -102.2421
6.0425 | -102.2271
6.0462 | -102.2108
6.0499 | -102.1876
6.0535 | -102.162
6.0572 | -102.1416
6.0609 | -102.1244
6.0646 | -102.1048
6.0682 | -102.0849
6.0719 | -102.0696
6.0756 | -102.0555
6.0793 | -102.0443
6.0829 | -102.0336
6.0866 | -102.0232
6.0903 | -102.0049
6.0939 | -101.986
6.0976 | -101.9626
6.1013 | -101.9331
6.105 | -101.9073
6.1086 | -101.8812
6.1123 | -101.8648
6.116 | -101.8549
6.1197 | -101.8496
6.1233 | -101.8523
6.127 | -101.8659
6.1307 | -101.8852
6.1344 | -101.9162
6.138 | -101.9638
6.1417 | -102.0197
6.1454 | -102.0838
6.149 | -102.1529
6.1527 | -102.2209
6.1564 | -102.2927
6.1601 | -102.3529
6.1637 | -102.3996
6.1674 | -102.4357
6.1711 | -102.4539
6.1748 | -102.4522
6.1784 | -102.4376
6.1821 | -102.4132
6.1858 | -102.3789
6.1894 | -102.3412
6.1931 | -102.3043
6.1968 | -102.2677
6.2005 | -102.2415
6.2041 | -102.2175
6.2078 | -102.2037
6.2115 | -102.1941
6.2152 | -102.1905
6.2188 | -102.1969
6.2225 | -102.2056
6.2262 | -102.2212
6.2299 | -102.2319
6.2335 | -102.2473
6.2372 | -102.2602
6.2409 | -102.2683
6.2445 | -102.2772
6.2482 | -102.2731
6.2519 | -102.2715
6.2556 | -102.2651
6.2592 | -102.2495
6.2629 | -102.2298
6.2666 | -102.2094
6.2703 | -102.1862
6.2739 | -102.1608
6.2776 | -102.1381
6.2813 | -102.1242
6.285 | -102.1202
6.2886 | -102.1259
6.2923 | -102.1423
6.296 | -102.1712
6.2996 | -102.2089
6.3033 | -102.2483
6.307 | -102.2826
6.3107 | -102.3122
6.3143 | -102.3301
6.318 | -102.3313
6.3217 | -102.322
6.3254 | -102.3071
6.329 | -102.2851
6.3327 | -102.2538
6.3364 | -102.2197
6.3401 | -102.1909
6.3437 | -102.1572
6.3474 | -102.1224
6.3511 | -102.0884
6.3547 | -102.0582
6.3584 | -102.0397
6.3621 | -102.0258
6.3658 | -102.0129
6.3694 | -102.0048
6.3731 | -102.0083
6.3768 | -102.0215
6.3805 | -102.0453
6.3841 | -102.0768
6.3878 | -102.1168
6.3915 | -102.1624
6.3952 | -102.2108
6.3988 | -102.2642
6.4025 | -102.3127
6.4062 | -102.3589
6.4098 | -102.4037
6.4135 | -102.4433
6.4172 | -102.4821
6.4209 | -102.5114
6.4245 | -102.5356
6.4282 | -102.563
6.4319 | -102.5853
6.4356 | -102.6048
6.4392 | -102.617
6.4429 | -102.6183
6.4466 | -102.6074
6.4502 | -102.5868
6.4539 | -102.5581
6.4576 | -102.5225
6.4613 | -102.48
6.4649 | -102.4394
6.4686 | -102.3958
6.4723 | -102.351
6.476 | -102.3176
6.4796 | -102.2959
6.4833 | -102.2865
6.487 | -102.2852
6.4907 | -102.2961
6.4943 | -102.3172
6.498 | -102.3463
6.5017 | -102.3818
6.5053 | -102.4204
6.509 | -102.464
6.5127 | -102.5064
6.5164 | -102.5419
6.52 | -102.5781
6.5237 | -102.6111
6.5274 | -102.6375
6.5311 | -102.6614
6.5347 | -102.6812
6.5384 | -102.6988
6.5421 | -102.7172
6.5458 | -102.736
6.5494 | -102.7488
6.5531 | -102.7635
6.5568 | -102.7797
6.5604 | -102.7869
6.5641 | -102.8011
6.5678 | -102.8211
6.5715 | -102.8316
6.5751 | -102.8259
6.5788 | -102.8125
6.5825 | -102.7977
6.5862 | -102.7775
6.5898 | -102.7434
6.5935 | -102.7022
6.5972 | -102.6571
6.6009 | -102.6111
6.6045 | -102.5642
6.6082 | -102.5204
6.6119 | -102.4848
6.6155 | -102.4559
6.6192 | -102.4364
6.6229 | -102.4273
6.6266 | -102.4291
6.6302 | -102.4336
6.6339 | -102.4495
6.6376 | -102.472
6.6413 | -102.4995
6.6449 | -102.5317
6.6486 | -102.5667
6.6523 | -102.596
6.656 | -102.6143
6.6596 | -102.6391
6.6633 | -102.6523
6.667 | -102.6576
6.6706 | -102.6585
6.6743 | -102.6624
6.678 | -102.6594
6.6817 | -102.6463
6.6853 | -102.6365
6.689 | -102.6337
6.6927 | -102.627
6.6964 | -102.6103
6.7 | -102.59
6.7037 | -102.5637
6.7074 | -102.5306
6.7111 | -102.4887
6.7147 | -102.4473
6.7184 | -102.4052
6.7221 | -102.3645
6.7257 | -102.3283
6.7294 | -102.3008
6.7331 | -102.2742
6.7368 | -102.2567
6.7404 | -102.247
6.7441 | -102.2399
6.7478 | -102.2413
6.7515 | -102.2502
6.7551 | -102.2667
6.7588 | -102.2819
6.7625 | -102.3001
6.7662 | -102.3215
6.7698 | -102.3517
6.7735 | -102.3849
6.7772 | -102.4292
6.7808 | -102.4926
6.7845 | -102.5603
6.7882 | -102.6343
6.7919 | -102.7208
6.7955 | -102.8145
6.7992 | -102.9043
6.8029 | -102.9934
6.8066 | -103.0766
6.8102 | -103.1449
6.8139 | -103.197
6.8176 | -103.237
6.8213 | -103.2635
6.8249 | -103.2747
6.8286 | -103.2707
6.8323 | -103.2559
6.8359 | -103.231
6.8396 | -103.1972
6.8433 | -103.1597
6.847 | -103.1136
6.8506 | -103.0679
6.8543 | -103.0232
6.858 | -102.9837
6.8617 | -102.9497
6.8653 | -102.9209
6.869 | -102.8927
6.8727 | -102.862
6.8763 | -102.8314
6.88 | -102.8004
6.8837 | -102.7732
6.8874 | -102.7437
6.891 | -102.7203
6.8947 | -102.7019
6.8984 | -102.685
6.9021 | -102.6705
6.9057 | -102.656
6.9094 | -102.644
6.9131 | -102.6353
6.9168 | -102.6246
6.9204 | -102.6122
6.9241 | -102.6025
6.9278 | -102.5947
6.9314 | -102.5869
6.9351 | -102.576
6.9388 | -102.5624
6.9425 | -102.5482
6.9461 | -102.5321
6.9498 | -102.5104
6.9535 | -102.4849
6.9572 | -102.462
6.9608 | -102.4408
6.9645 | -102.4221
6.9682 | -102.4078
6.9719 | -102.4029
6.9755 | -102.407
6.9792 | -102.4219
6.9829 | -102.4399
6.9865 | -102.4672
6.9902 | -102.5025
6.9939 | -102.541
6.9976 | -102.5798
7.0012 | -102.6181
7.0049 | -102.6599
7.0086 | -102.7005
7.0123 | -102.7407
7.0159 | -102.7785
7.0196 | -102.8111
7.0233 | -102.8409
7.027 | -102.8688
7.0306 | -102.8923
7.0343 | -102.9131
7.038 | -102.9327
7.0416 | -102.9489
7.0453 | -102.9616
7.049 | -102.9654
7.0527 | -102.9656
7.0563 | -102.97
7.06 | -102.9742
7.0637 | -102.9706
7.0674 | -102.9643
7.071 | -102.9592
7.0747 | -102.9518
7.0784 | -102.9466
7.0821 | -102.9434
7.0857 | -102.9335
7.0894 | -102.931
7.0931 | -102.9313
7.0967 | -102.9321
7.1004 | -102.9354
7.1041 | -102.9426
7.1078 | -102.9521
7.1114 | -102.967
7.1151 | -102.9857
7.1188 | -103.0028
7.1225 | -103.0202
7.1261 | -103.0373
7.1298 | -103.0439
7.1335 | -103.0497
7.1372 | -103.052
7.1408 | -103.0463
7.1445 | -103.0394
7.1482 | -103.0275
7.1518 | -103.0128
7.1555 | -103.0018
7.1592 | -102.9877
7.1629 | -102.9732
7.1665 | -102.9618
7.1702 | -102.9585
7.1739 | -102.9645
7.1776 | -102.9797
7.1812 | -103.0035
7.1849 | -103.0264
7.1886 | -103.0545
7.1922 | -103.0912
7.1959 | -103.1257
7.1996 | -103.1566
7.2033 | -103.189
7.2069 | -103.2165
7.2106 | -103.2359
7.2143 | -103.2458
7.218 | -103.2459
7.2216 | -103.2419
7.2253 | -103.2365
7.229 | -103.2303
7.2327 | -103.2275
7.2363 | -103.2271
7.24 | -103.2201
7.2437 | -103.2145
7.2473 | -103.2114
7.251 | -103.2035
7.2547 | -103.1998
7.2584 | -103.1963
7.262 | -103.1876
7.2657 | -103.1743
7.2694 | -103.155
7.2731 | -103.1381
7.2767 | -103.1207
7.2804 | -103.1035
7.2841 | -103.0846
7.2878 | -103.0706
7.2914 | -103.069
7.2951 | -103.0709
7.2988 | -103.071
7.3024 | -103.0749
7.3061 | -103.084
7.3098 | -103.0949
7.3135 | -103.1088
7.3171 | -103.1244
7.3208 | -103.1401
7.3245 | -103.1537
7.3282 | -103.1591
7.3318 | -103.1553
7.3355 | -103.1443
7.3392 | -103.1263
7.3429 | -103.1073
7.3465 | -103.0857
7.3502 | -103.0682
7.3539 | -103.0566
7.3575 | -103.0555
7.3612 | -103.0591
7.3649 | -103.0726
7.3686 | -103.0932
7.3722 | -103.1242
7.3759 | -103.165
7.3796 | -103.2062
7.3833 | -103.2537
7.3869 | -103.3035
7.3906 | -103.358
7.3943 | -103.4039
7.398 | -103.4422
7.4016 | -103.4837
7.4053 | -103.521
7.409 | -103.5472
7.4126 | -103.5686
7.4163 | -103.5603
7.42 | -103.5612
7.4237 | -103.5599
7.4273 | -103.5518
7.431 | -103.5392
7.4347 | -103.5196
7.4384 | -103.4977
7.442 | -103.4722
7.4457 | -103.4425
7.4494 | -103.4114
7.4531 | -103.3776
7.4567 | -103.3439
7.4604 | -103.3099
7.4641 | -103.2802
7.4677 | -103.2516
7.4714 | -103.2281
7.4751 | -103.2068
7.4788 | -103.1942
7.4824 | -103.1868
7.4861 | -103.1796
7.4898 | -103.1739
7.4935 | -103.1762
7.4971 | -103.185
7.5008 | -103.1929
7.5045 | -103.2061
7.5081 | -103.2206
7.5118 | -103.231
7.5155 | -103.2483
7.5192 | -103.263
7.5228 | -103.2741
7.5265 | -103.2848
7.5302 | -103.3008
7.5339 | -103.3142
7.5375 | -103.3225
7.5412 | -103.3315
7.5449 | -103.3433
7.5486 | -103.357
7.5522 | -103.3695
7.5559 | -103.3834
7.5596 | -103.3999
7.5632 | -103.4186
7.5669 | -103.4396
7.5706 | -103.4647
7.5743 | -103.4873
7.5779 | -103.4985
7.5816 | -103.516
7.5853 | -103.5336
7.589 | -103.5482
7.5926 | -103.5572
7.5963 | -103.5599
7.6 | -103.5559
7.6037 | -103.552
7.6073 | -103.5397
7.611 | -103.518
7.6147 | -103.5013
7.6183 | -103.4874
7.622 | -103.4721
7.6257 | -103.4613
7.6294 | -103.4487
7.633 | -103.4466
7.6367 | -103.4494
7.6404 | -103.4562
7.6441 | -103.4672
7.6477 | -103.484
7.6514 | -103.5075
7.6551 | -103.5329
7.6588 | -103.562
7.6624 | -103.5984
7.6661 | -103.6381
7.6698 | -103.6751
7.6734 | -103.704
7.6771 | -103.7267
7.6808 | -103.7466
7.6845 | -103.7576
7.6881 | -103.7603
7.6918 | -103.7552
7.6955 | -103.7313
7.6992 | -103.7019
7.7028 | -103.6627
7.7065 | -103.6125
7.7102 | -103.5617
7.7139 | -103.5086
7.7175 | -103.4458
7.7212 | -103.3846
7.7249 | -103.3241
7.7285 | -103.264
7.7322 | -103.2072
7.7359 | -103.1519
7.7396 | -103.1066
7.7432 | -103.0714
7.7469 | -103.0428
7.7506 | -103.0234
7.7543 | -103.0148
7.7579 | -103.0196
7.7616 | -103.0347
7.7653 | -103.0571
7.769 | -103.0908
7.7726 | -103.126
7.7763 | -103.1657
7.78 | -103.2095
7.7836 | -103.2568
7.7873 | -103.3045
7.791 | -103.3491
7.7947 | -103.3904
7.7983 | -103.4331
7.802 | -103.4742
7.8057 | -103.5206
7.8094 | -103.5659
7.813 | -103.6058
7.8167 | -103.6498
7.8204 | -103.6992
7.8241 | -103.7405
7.8277 | -103.7846
7.8314 | -103.8335
7.8351 | -103.8768
7.8387 | -103.9185
7.8424 | -103.9536
7.8461 | -103.9763
7.8498 | -103.9958
7.8534 | -104.0061
7.8571 | -104.0081
7.8608 | -104.0015
7.8645 | -103.9903
7.8681 | -103.9754
7.8718 | -103.9541
7.8755 | -103.9344
7.8792 | -103.9181
7.8828 | -103.9036
7.8865 | -103.8888
7.8902 | -103.8661
7.8938 | -103.8446
7.8975 | -103.8226
7.9012 | -103.8088
7.9049 | -103.7959
7.9085 | -103.7801
7.9122 | -103.7647
7.9159 | -103.7502
7.9196 | -103.7321
7.9232 | -103.7166
7.9269 | -103.7045
7.9306 | -103.6914
7.9342 | -103.6775
7.9379 | -103.6691
7.9416 | -103.6579
7.9453 | -103.6516
7.9489 | -103.67
7.9526 | -103.6763
7.9563 | -103.6846
7.96 | -103.693
7.9636 | -103.7106
7.9673 | -103.7242
7.971 | -103.7352
7.9747 | -103.7498
7.9783 | -103.7604
7.982 | -103.7601
7.9857 | -103.7582
7.9893 | -103.7592
7.993 | -103.7535
7.9967 | -103.7426
8.0004 | -103.7301
8.004 | -103.7205
8.0077 | -103.7136
8.0114 | -103.7054
8.0151 | -103.6965
8.0187 | -103.6845
8.0224 | -103.6735
8.0261 | -103.6646
8.0298 | -103.6585
8.0334 | -103.6569
8.0371 | -103.6625
8.0408 | -103.6688
8.0444 | -103.6773
8.0481 | -103.6857
8.0518 | -103.6926
8.0555 | -103.6955
8.0591 | -103.6963
8.0628 | -103.694
8.0665 | -103.6881
8.0702 | -103.6755
8.0738 | -103.662
8.0775 | -103.6468
8.0812 | -103.6302
8.0849 | -103.6031
8.0885 | -103.5803
8.0922 | -103.5598
8.0959 | -103.5352
8.0995 | -103.5149
8.1032 | -103.4965
8.1069 | -103.4764
8.1106 | -103.4587
8.1142 | -103.4411
8.1179 | -103.4234
8.1216 | -103.4073
8.1253 | -103.3948
8.1289 | -103.3831
8.1326 | -103.3724
8.1363 | -103.3563
8.14 | -103.3366
8.1436 | -103.3186
8.1473 | -103.3007
8.151 | -103.2791
8.1546 | -103.2495
8.1583 | -103.2175
8.162 | -103.1911
8.1657 | -103.1673
8.1693 | -103.1513
8.173 | -103.1368
8.1767 | -103.1268
8.1804 | -103.1248
8.184 | -103.124
8.1877 | -103.1274
8.1914 | -103.137
8.1951 | -103.1536
8.1987 | -103.1757
8.2024 | -103.203
8.2061 | -103.2433
8.2097 | -103.2834
8.2134 | -103.3262
8.2171 | -103.3745
8.2208 | -103.424
8.2244 | -103.4795
8.2281 | -103.5346
8.2318 | -103.5829
8.2355 | -103.6298
8.2391 | -103.6737
8.2428 | -103.7093
8.2465 | -103.7385
8.2501 | -103.7572
8.2538 | -103.7713
8.2575 | -103.7851
8.2612 | -103.7907
8.2648 | -103.7893
8.2685 | -103.7837
8.2722 | -103.7752
8.2759 | -103.7619
8.2795 | -103.7447
8.2832 | -103.7262
8.2869 | -103.7116
8.2906 | -103.6956
8.2942 | -103.6854
8.2979 | -103.6761
8.3016 | -103.6722
8.3052 | -103.6764
8.3089 | -103.6835
8.3126 | -103.6939
8.3163 | -103.7068
8.3199 | -103.7243
8.3236 | -103.744
8.3273 | -103.7577
8.331 | -103.7745
8.3346 | -103.7924
8.3383 | -103.8095
8.342 | -103.8202
8.3457 | -103.8278
8.3493 | -103.8283
8.353 | -103.8262
8.3567 | -103.8259
8.3603 | -103.8208
8.364 | -103.8113
8.3677 | -103.7943
8.3714 | -103.773
8.375 | -103.7503
8.3787 | -103.7233
8.3824 | -103.6993
8.3861 | -103.6762
8.3897 | -103.6543
8.3934 | -103.6354
8.3971 | -103.6251
8.4008 | -103.6131
8.4044 | -103.6017
8.4081 | -103.5956
8.4118 | -103.6017
8.4154 | -103.616
8.4191 | -103.6376
8.4228 | -103.6586
8.4265 | -103.686
8.4301 | -103.7222
8.4338 | -103.763
8.4375 | -103.8021
8.4412 | -103.8394
8.4448 | -103.8752
8.4485 | -103.9128
8.4522 | -103.9521
8.4559 | -103.9848
8.4595 | -104.0133
8.4632 | -104.0408
8.4669 | -104.0616
8.4705 | -104.0797
8.4742 | -104.0888
8.4779 | -104.0735
8.4816 | -104.0683
8.4852 | -104.0594
8.4889 | -104.0513
8.4926 | -104.0443
8.4963 | -104.0426
8.4999 | -104.0342
8.5036 | -104.0277
8.5073 | -104.021
8.511 | -104.0203
8.5146 | -104.0302
8.5183 | -104.0456
8.522 | -104.0646
8.5256 | -104.0925
8.5293 | -104.1273
8.533 | -104.1622
8.5367 | -104.1929
8.5403 | -104.2267
8.544 | -104.2567
8.5477 | -104.2911
8.5514 | -104.322
8.555 | -104.3463
8.5587 | -104.3652
8.5624 | -104.3748
8.5661 | -104.3784
8.5697 | -104.3754
8.5734 | -104.363
8.5771 | -104.3423
8.5807 | -104.3166
8.5844 | -104.2863
8.5881 | -104.2529
8.5918 | -104.2208
8.5954 | -104.1897
8.5991 | -104.1567
8.6028 | -104.1293
8.6065 | -104.1035
8.6101 | -104.0816
8.6138 | -104.062
8.6175 | -104.0504
8.6211 | -104.0421
8.6248 | -104.0396
8.6285 | -104.0473
8.6322 | -104.0617
8.6358 | -104.0809
8.6395 | -104.1034
8.6432 | -104.1276
8.6469 | -104.1507
8.6505 | -104.1802
8.6542 | -104.2069
8.6579 | -104.2307
8.6616 | -104.2502
8.6652 | -104.2688
8.6689 | -104.2796
8.6726 | -104.2851
8.6762 | -104.2812
8.6799 | -104.2727
8.6836 | -104.2636
8.6873 | -104.2487
8.6909 | -104.2301
8.6946 | -104.207
8.6983 | -104.1775
8.702 | -104.1464
8.7056 | -104.1188
8.7093 | -104.0882
8.713 | -104.0595
8.7167 | -104.035
8.7203 | -104.0109
8.724 | -103.9886
8.7277 | -103.9774
8.7313 | -103.9673
8.735 | -103.9592
8.7387 | -103.9605
8.7424 | -103.9674
8.746 | -103.9708
8.7497 | -103.9791
8.7534 | -103.9887
8.7571 | -103.9987
8.7607 | -104.0067
8.7644 | -104.0157
8.7681 | -104.0243
8.7718 | -104.0292
8.7754 | -104.0373
8.7791 | -104.0436
8.7828 | -104.0515
8.7864 | -104.0584
8.7901 | -104.0681
8.7938 | -104.0797
8.7975 | -104.0948
8.8011 | -104.1119
8.8048 | -104.1281
8.8085 | -104.1443
8.8122 | -104.1624
8.8158 | -104.1859
8.8195 | -104.2119
8.8232 | -104.2363
8.8269 | -104.2658
8.8305 | -104.2962
8.8342 | -104.3224
8.8379 | -104.3544
8.8415 | -104.3823
8.8452 | -104.4057
8.8489 | -104.4328
8.8526 | -104.4609
8.8562 | -104.4857
8.8599 | -104.5034
8.8636 | -104.5161
8.8673 | -104.5199
8.8709 | -104.5156
8.8746 | -104.5036
8.8783 | -104.4806
8.882 | -104.4488
8.8856 | -104.4122
8.8893 | -104.3721
8.893 | -104.3273
8.8966 | -104.2816
8.9003 | -104.2345
8.904 | -104.1839
8.9077 | -104.1373
8.9113 | -104.0935
8.915 | -104.0585
8.9187 | -104.0311
8.9224 | -104.0032
8.926 | -103.9821
8.9297 | -103.9699
8.9334 | -103.9664
8.937 | -103.9675
8.9407 | -103.971
8.9444 | -103.98
8.9481 | -103.9927
8.9517 | -104.0096
8.9554 | -104.0274
8.9591 | -104.0465
8.9628 | -104.0686
8.9664 | -104.0894
8.9701 | -104.1114
8.9738 | -104.1338
8.9775 | -104.15
8.9811 | -104.1678
8.9848 | -104.1797
8.9885 | -104.1929
8.9921 | -104.2076
8.9958 | -104.2176
8.9995 | -104.2264
9.0032 | -104.236
9.0068 | -104.2384
9.0105 | -104.2457
9.0142 | -104.2454
9.0179 | -104.2382
9.0215 | -104.2278
9.0252 | -104.2155
9.0289 | -104.2009
9.0326 | -104.1867
9.0362 | -104.1747
9.0399 | -104.1562
9.0436 | -104.1361
9.0472 | -104.1149
9.0509 | -104.0893
9.0546 | -104.0677
9.0583 | -104.0479
9.0619 | -104.0301
9.0656 | -104.0121
9.0693 | -103.9943
9.073 | -103.9763
9.0766 | -103.9625
9.0803 | -103.9508
9.084 | -103.9448
9.0877 | -103.9435
9.0913 | -103.9443
9.095 | -103.9456
9.0987 | -103.9529
9.1023 | -103.9665
9.106 | -103.9797
9.1097 | -103.9949
9.1134 | -104.0099
9.117 | -104.0251
9.1207 | -104.0421
9.1244 | -104.056
9.1281 | -104.0707
9.1317 | -104.0834
9.1354 | -104.0931
9.1391 | -104.1018
9.1428 | -104.1121
9.1464 | -104.1239
9.1501 | -104.1375
9.1538 | -104.1495
9.1574 | -104.1624
9.1611 | -104.1749
9.1648 | -104.1842
9.1685 | -104.1998
9.1721 | -104.2169
9.1758 | -104.2353
9.1795 | -104.2566
9.1832 | -104.2731
9.1868 | -104.296
9.1905 | -104.3171
9.1942 | -104.3337
9.1979 | -104.3454
9.2015 | -104.3567
9.2052 | -104.365
9.2089 | -104.3696
9.2125 | -104.3681
9.2162 | -104.3642
9.2199 | -104.3529
9.2236 | -104.3414
9.2272 | -104.3312
9.2309 | -104.3153
9.2346 | -104.2979
9.2383 | -104.2764
9.2419 | -104.2554
9.2456 | -104.2377
9.2493 | -104.2155
9.253 | -104.1984
9.2566 | -104.1853
9.2603 | -104.1747
9.264 | -104.1663
9.2676 | -104.1609
9.2713 | -104.1613
9.275 | -104.1632
9.2787 | -104.1653
9.2823 | -104.1713
9.286 | -104.18
9.2897 | -104.1935
9.2934 | -104.2041
9.297 | -104.2152
9.3007 | -104.2308
9.3044 | -104.2449
9.308 | -104.2584
9.3117 | -104.2735
9.3154 | -104.2885
9.3191 | -104.3084
9.3227 | -104.3256
9.3264 | -104.3432
9.3301 | -104.359
9.3338 | -104.3765
9.3374 | -104.391
9.3411 | -104.4082
9.3448 | -104.4263
9.3485 | -104.4403
9.3521 | -104.453
9.3558 | -104.4608
9.3595 | -104.4649
9.3631 | -104.4685
9.3668 | -104.4775
9.3705 | -104.4871
9.3742 | -104.4966
9.3778 | -104.5102
9.3815 | -104.5222
9.3852 | -104.5458
9.3889 | -104.565
9.3925 | -104.5848
9.3962 | -104.6066
9.3999 | -104.6291
9.4036 | -104.6574
9.4072 | -104.6819
9.4109 | -104.6999
9.4146 | -104.7107
9.4182 | -104.7176
9.4219 | -104.7247
9.4256 | -104.7224
9.4293 | -104.7169
9.4329 | -104.7093
9.4366 | -104.698
9.4403 | -104.6794
9.444 | -104.66
9.4476 | -104.6419
9.4513 | -104.6235
9.455 | -104.6057
9.4587 | -104.592
9.4623 | -104.5817
9.466 | -104.5762
9.4697 | -104.5764
9.4733 | -104.5834
9.477 | -104.5956
9.4807 | -104.6111
9.4844 | -104.6354
9.488 | -104.6638
9.4917 | -104.6973
9.4954 | -104.7324
9.4991 | -104.7688
9.5027 | -104.8018
9.5064 | -104.8313
9.5101 | -104.8598
9.5138 | -104.882
9.5174 | -104.9049
9.5211 | -104.9198
9.5248 | -104.9299
9.5284 | -104.9304
9.5321 | -104.9323
9.5358 | -104.9225
9.5395 | -104.9223
9.5431 | -104.9147
9.5468 | -104.9078
9.5505 | -104.9047
9.5542 | -104.9034
9.5578 | -104.8966
9.5615 | -104.8861
9.5652 | -104.8775
9.5689 | -104.8662
9.5725 | -104.8525
9.5762 | -104.8424
9.5799 | -104.837
9.5835 | -104.8342
9.5872 | -104.8318
9.5909 | -104.8343
9.5946 | -104.8363
9.5982 | -104.8364
9.6019 | -104.8396
9.6056 | -104.8429
9.6093 | -104.8464
9.6129 | -104.8497
9.6166 | -104.8481
9.6203 | -104.8419
9.624 | -104.8308
9.6276 | -104.8235
9.6313 | -104.8183
9.635 | -104.8084
9.6386 | -104.7953
9.6423 | -104.7744
9.646 | -104.7533
9.6497 | -104.7307
9.6533 | -104.7065
9.657 | -104.6827
9.6607 | -104.6551
9.6644 | -104.6294
9.668 | -104.6032
9.6717 | -104.5714
9.6754 | -104.5391
9.679 | -104.509
9.6827 | -104.4811
9.6864 | -104.4559
9.6901 | -104.4342
9.6937 | -104.4105
9.6974 | -104.3881
9.7011 | -104.3704
9.7048 | -104.355
9.7084 | -104.3432
9.7121 | -104.3345
9.7158 | -104.3344
9.7195 | -104.3342
9.7231 | -104.3342
9.7268 | -104.3374
9.7305 | -104.344
9.7341 | -104.3488
9.7378 | -104.356
9.7415 | -104.3658
9.7452 | -104.3712
9.7488 | -104.3775
9.7525 | -104.3826
9.7562 | -104.3888
9.7599 | -104.4002
9.7635 | -104.4143
9.7672 | -104.4309
9.7709 | -104.4447
9.7746 | -104.4583
9.7782 | -104.4784
9.7819 | -104.5035
9.7856 | -104.5363
9.7892 | -104.5665
9.7929 | -104.6007
9.7966 | -104.6364
9.8003 | -104.6706
9.8039 | -104.7104
9.8076 | -104.7504
9.8113 | -104.7909
9.815 | -104.8288
9.8186 | -104.8603
9.8223 | -104.8845
9.826 | -104.8997
9.8297 | -104.9094
9.8333 | -104.9155
9.837 | -104.9127
9.8407 | -104.9005
9.8443 | -104.8816
9.848 | -104.8609
9.8517 | -104.8346
9.8554 | -104.8031
9.859 | -104.7611
9.8627 | -104.7171
9.8664 | -104.6756
9.8701 | -104.6339
9.8737 | -104.5926
9.8774 | -104.5543
9.8811 | -104.5188
9.8848 | -104.4883
9.8884 | -104.4643
9.8921 | -104.4455
9.8958 | -104.433
9.8994 | -104.4225
9.9031 | -104.4197
9.9068 | -104.4232
9.9105 | -104.4316
9.9141 | -104.4443
9.9178 | -104.4593
9.9215 | -104.4769
9.9252 | -104.4941
9.9288 | -104.5141
9.9325 | -104.5358
9.9362 | -104.5572
9.9398 | -104.5778
9.9435 | -104.5935
9.9472 | -104.6083
9.9509 | -104.62
9.9545 | -104.6306
9.9582 | -104.6413
9.9619 | -104.6516
9.9656 | -104.6626
9.9692 | -104.6728
9.9729 | -104.6807
9.9766 | -104.6844
9.9803 | -104.6929
9.9839 | -104.7035
9.9876 | -104.7047
9.9913 | -104.7017
9.9949 | -104.7044
9.9986 | -104.7101
10.0023 | -104.719
10.006 | -104.7304
10.0096 | -104.7437
10.0133 | -104.7565
10.017 | -104.7712
10.0207 | -104.7886
10.0243 | -104.8029
10.028 | -104.8099
10.0317 | -104.8189
10.0354 | -104.828
10.039 | -104.833
10.0427 | -104.8441
10.0464 | -104.8525
10.05 | -104.8607
10.0537 | -104.8693
10.0574 | -104.8793
10.0611 | -104.8936
10.0647 | -104.9102
10.0684 | -104.9272
10.0721 | -104.9422
10.0758 | -104.9568
10.0794 | -104.9727
10.0831 | -104.9828
10.0868 | -104.9943
10.0905 | -105.0064
10.0941 | -105.0111
10.0978 | -105.0126
10.1015 | -105.0103
10.1051 | -105.0085
10.1088 | -105.0025
10.1125 | -104.9912
10.1162 | -104.9792
10.1198 | -104.9642
10.1235 | -104.949
10.1272 | -104.9319
10.1309 | -104.9163
10.1345 | -104.9079
10.1382 | -104.8984
10.1419 | -104.8831
10.1456 | -104.8704
10.1492 | -104.8589
10.1529 | -104.8517
10.1566 | -104.8462
10.1602 | -104.8443
10.1639 | -104.8423
10.1676 | -104.8418
10.1713 | -104.8439
10.1749 | -104.8502
10.1786 | -104.8555
10.1823 | -104.8584
10.186 | -104.861
10.1896 | -104.8588
10.1933 | -104.8559
10.197 | -104.8547
10.2007 | -104.8489
10.2043 | -104.8416
10.208 | -104.832
10.2117 | -104.8163
10.2153 | -104.8037
10.219 | -104.787
10.2227 | -104.7708
10.2264 | -104.7591
10.23 | -104.7474
10.2337 | -104.7395
10.2374 | -104.7325
10.2411 | -104.7282
10.2447 | -104.7274
10.2484 | -104.7296
10.2521 | -104.7371
10.2558 | -104.7447
10.2594 | -104.7588
10.2631 | -104.7792
10.2668 | -104.8011
10.2704 | -104.8261
10.2741 | -104.8519
10.2778 | -104.8775
10.2815 | -104.9025
10.2851 | -104.9277
10.2888 | -104.9507
10.2925 | -104.975
10.2962 | -104.9935
10.2998 | -105.0066
10.3035 | -105.0137
10.3072 | -105.0142
10.3109 | -105.0141
10.3145 | -105.0134
10.3182 | -105.0102
10.3219 | -105.0035
10.3255 | -104.9877
10.3292 | -104.9661
10.3329 | -104.9471
10.3366 | -104.927
10.3402 | -104.9028
10.3439 | -104.8779
10.3476 | -104.8491
10.3513 | -104.8226
10.3549 | -104.7958
10.3586 | -104.7707
10.3623 | -104.7486
10.3659 | -104.7295
10.3696 | -104.7139
10.3733 | -104.7021
10.377 | -104.6909
10.3806 | -104.6842
10.3843 | -104.6804
10.388 | -104.6806
10.3917 | -104.6842
10.3953 | -104.6894
10.399 | -104.6978
10.4027 | -104.7046
10.4064 | -104.7168
10.41 | -104.7327
10.4137 | -104.7453
10.4174 | -104.7602
10.421 | -104.7769
10.4247 | -104.7947
10.4284 | -104.8125
10.4321 | -104.8291
10.4357 | -104.8441
10.4394 | -104.8583
10.4431 | -104.8667
10.4468 | -104.8748
10.4504 | -104.8789
10.4541 | -104.8794
10.4578 | -104.8783
10.4615 | -104.8691
10.4651 | -104.854
10.4688 | -104.8364
10.4725 | -104.8179
10.4761 | -104.8038
10.4798 | -104.7896
10.4835 | -104.7734
10.4872 | -104.7611
10.4908 | -104.7502
10.4945 | -104.7377
10.4982 | -104.7305
10.5019 | -104.7234
10.5055 | -104.7194
10.5092 | -104.7176
10.5129 | -104.7206
10.5166 | -104.7253
10.5202 | -104.7333
10.5239 | -104.7445
10.5276 | -104.7568
10.5312 | -104.7712
10.5349 | -104.7909
10.5386 | -104.8124
10.5423 | -104.8361
10.5459 | -104.8654
10.5496 | -104.8935
10.5533 | -104.9227
10.557 | -104.954
10.5606 | -104.9847
10.5643 | -105.0186
10.568 | -105.0513
10.5717 | -105.0852
10.5753 | -105.1203
10.579 | -105.1511
10.5827 | -105.1729
10.5863 | -105.1945
10.59 | -105.2181
10.5937 | -105.2402
10.5974 | -105.2461
10.601 | -105.2576
10.6047 | -105.2669
10.6084 | -105.2731
10.6121 | -105.2742
10.6157 | -105.2709
10.6194 | -105.2651
10.6231 | -105.2555
10.6268 | -105.2486
10.6304 | -105.2357
10.6341 | -105.2195
10.6378 | -105.2016
10.6414 | -105.1794
10.6451 | -105.1603
10.6488 | -105.1445
10.6525 | -105.1237
10.6561 | -105.1051
10.6598 | -105.0853
10.6635 | -105.0647
10.6672 | -105.0425
10.6708 | -105.0186
10.6745 | -104.9975
10.6782 | -104.9797
10.6818 | -104.9629
10.6855 | -104.9464
10.6892 | -104.9291
10.6929 | -104.9133
10.6965 | -104.9009
10.7002 | -104.8859
10.7039 | -104.8736
10.7076 | -104.8598
10.7112 | -104.8454
10.7149 | -104.8328
10.7186 | -104.8238
10.7223 | -104.8144
10.7259 | -104.8028
10.7296 | -104.7877
10.7333 | -104.7755
10.7369 | -104.7683
10.7406 | -104.7581
10.7443 | -104.7436
10.748 | -104.7302
10.7516 | -104.7171
10.7553 | -104.7061
10.759 | -104.6993
10.7627 | -104.6929
10.7663 | -104.6875
10.77 | -104.6874
10.7737 | -104.6938
10.7774 | -104.7024
10.781 | -104.7113
10.7847 | -104.719
10.7884 | -104.7315
10.792 | -104.7491
10.7957 | -104.7696
10.7994 | -104.7927
10.8031 | -104.813
10.8067 | -104.829
10.8104 | -104.8486
10.8141 | -104.8649
10.8178 | -104.8802
10.8214 | -104.8941
10.8251 | -104.9053
10.8288 | -104.9132
10.8325 | -104.9168
10.8361 | -104.9157
10.8398 | -104.9146
10.8435 | -104.9137
10.8471 | -104.9132
10.8508 | -104.9103
10.8545 | -104.9017
10.8582 | -104.8948
10.8618 | -104.8903
10.8655 | -104.8921
10.8692 | -104.8899
10.8729 | -104.8888
10.8765 | -104.8896
10.8802 | -104.8956
10.8839 | -104.9008
10.8876 | -104.9051
10.8912 | -104.9104
10.8949 | -104.9138
10.8986 | -104.9177
10.9022 | -104.9258
10.9059 | -104.9318
10.9096 | -104.9396
10.9133 | -104.9467
10.9169 | -104.9526
10.9206 | -104.9599
10.9243 | -104.9652
10.928 | -104.9674
10.9316 | -104.9714
10.9353 | -104.9778
10.939 | -104.985
10.9427 | -104.9935
10.9463 | -105.0013
10.95 | -105.0108
10.9537 | -105.018
10.9573 | -105.0261
10.961 | -105.0383
10.9647 | -105.0547
10.9684 | -105.0679
10.972 | -105.0829
10.9757 | -105.0959
10.9794 | -105.1098
10.9831 | -105.1274
10.9867 | -105.1465
10.9904 | -105.1641
10.9941 | -105.1824
10.9977 | -105.1989
11.0014 | -105.2144
11.0051 | -105.2282
11.0088 | -105.2437
11.0124 | -105.261
11.0161 | -105.2777
11.0198 | -105.2931
11.0235 | -105.3056
11.0271 | -105.3193
11.0308 | -105.3322
11.0345 | -105.3405
11.0382 | -105.3461
11.0418 | -105.3522
11.0455 | -105.3629
11.0492 | -105.3691
11.0528 | -105.3717
11.0565 | -105.3639
11.0602 | -105.3498
11.0639 | -105.3335
11.0675 | -105.313
11.0712 | -105.2885
11.0749 | -105.2599
11.0786 | -105.2234
11.0822 | -105.1878
11.0859 | -105.1532
11.0896 | -105.1123
11.0933 | -105.0722
11.0969 | -105.0306
11.1006 | -104.9898
11.1043 | -104.952
11.1079 | -104.914
11.1116 | -104.881
11.1153 | -104.8521
11.119 | -104.827
11.1226 | -104.806
11.1263 | -104.8079
11.13 | -104.7957
11.1337 | -104.7833
11.1373 | -104.7747
11.141 | -104.7733
11.1447 | -104.7773
11.1484 | -104.7841
11.152 | -104.7944
11.1557 | -104.8043
11.1594 | -104.8158
11.163 | -104.8275
11.1667 | -104.8387
11.1704 | -104.8487
11.1741 | -104.857
11.1777 | -104.8628
11.1814 | -104.8646
11.1851 | -104.8659
11.1888 | -104.8702
11.1924 | -104.8704
11.1961 | -104.8692
11.1998 | -104.8668
11.2035 | -104.8644
11.2071 | -104.8597
11.2108 | -104.8567
11.2145 | -104.8544
11.2181 | -104.8565
11.2218 | -104.8611
11.2255 | -104.8667
11.2292 | -104.873
11.2328 | -104.8798
11.2365 | -104.8923
11.2402 | -104.9076
11.2439 | -104.922
11.2475 | -104.933
11.2512 | -104.9456
11.2549 | -104.9631
11.2586 | -104.9855
11.2622 | -105.0072
11.2659 | -105.0295
11.2696 | -105.0498
11.2732 | -105.0665
11.2769 | -105.0795
11.2806 | -105.088
11.2843 | -105.0953
11.2879 | -105.1036
11.2916 | -105.1072
11.2953 | -105.1087
11.299 | -105.1094
11.3026 | -105.109
11.3063 | -105.0987
11.31 | -105.0815
11.3137 | -105.0698
11.3173 | -105.0578
11.321 | -105.0429
11.3247 | -105.0304
11.3283 | -105.019
11.332 | -105.0039
11.3357 | -104.9932
11.3394 | -104.9839
11.343 | -104.9752
11.3467 | -104.9713
11.3504 | -104.9681
11.3541 | -104.9655
11.3577 | -104.9664
11.3614 | -104.9665
11.3651 | -104.9713
11.3688 | -104.9776
11.3724 | -104.9838
11.3761 | -104.9892
11.3798 | -104.9949
11.3834 | -104.9991
11.3871 | -104.9997
11.3908 | -105.0004
11.3945 | -105.0027
11.3981 | -105.0042
11.4018 | -105.0009
11.4055 | -104.9986
11.4092 | -104.9924
11.4128 | -104.9898
11.4165 | -104.9852
11.4202 | -104.977
11.4238 | -104.9662
11.4275 | -104.9526
11.4312 | -104.9405
11.4349 | -104.9293
11.4385 | -104.9199
11.4422 | -104.9138
11.4459 | -104.9055
11.4496 | -104.8934
11.4532 | -104.884
11.4569 | -104.877
11.4606 | -104.8718
11.4643 | -104.8707
11.4679 | -104.869
11.4716 | -104.8688
11.4753 | -104.872
11.4789 | -104.8749
11.4826 | -104.8783
11.4863 | -104.8861
11.49 | -104.8949
11.4936 | -104.9062
11.4973 | -104.9224
11.501 | -104.9332
11.5047 | -104.9452
11.5083 | -104.9604
11.512 | -104.9745
11.5157 | -104.9874
11.5194 | -104.9968
11.523 | -105.0074
11.5267 | -105.0203
11.5304 | -105.0293
11.534 | -105.0335
11.5377 | -105.038
11.5414 | -105.044
11.5451 | -105.0482
11.5487 | -105.0527
11.5524 | -105.0559
11.5561 | -105.0541
11.5598 | -105.0485
11.5634 | -105.043
11.5671 | -105.037
11.5708 | -105.0301
11.5745 | -105.0235
11.5781 | -105.0148
11.5818 | -105.0068
11.5855 | -104.9969
11.5891 | -104.992
11.5928 | -104.9922
11.5965 | -104.9915
11.6002 | -104.9938
11.6038 | -104.9956
11.6075 | -104.9984
11.6112 | -105.0075
11.6149 | -105.0191
11.6185 | -105.0336
11.6222 | -105.0474
11.6259 | -105.0648
11.6296 | -105.0811
11.6332 | -105.1036
11.6369 | -105.1316
11.6406 | -105.1589
11.6442 | -105.1829
11.6479 | -105.2083
11.6516 | -105.2333
11.6553 | -105.2464
11.6589 | -105.263
11.6626 | -105.2755
11.6663 | -105.287
11.67 | -105.3003
11.6736 | -105.3102
11.6773 | -105.3146
11.681 | -105.3135
11.6846 | -105.3089
11.6883 | -105.3035
11.692 | -105.2948
11.6957 | -105.2825
11.6993 | -105.2671
11.703 | -105.2528
11.7067 | -105.2306
11.7104 | -105.2049
11.714 | -105.1824
11.7177 | -105.1607
11.7214 | -105.1342
11.7251 | -105.1079
11.7287 | -105.0849
11.7324 | -105.0589
11.7361 | -105.036
11.7397 | -105.0146
11.7434 | -104.9941
11.7471 | -104.9751
11.7508 | -104.9582
11.7544 | -104.9431
11.7581 | -104.9295
11.7618 | -104.9207
11.7655 | -104.9148
11.7691 | -104.9103
11.7728 | -104.9133
11.7765 | -104.9173
11.7802 | -104.9244
11.7838 | -104.9346
11.7875 | -104.9467
11.7912 | -104.9665
11.7948 | -104.9841
11.7985 | -105.005
11.8022 | -105.0245
11.8059 | -105.0446
11.8095 | -105.0686
11.8132 | -105.0944
11.8169 | -105.1209
11.8206 | -105.1474
11.8242 | -105.1732
11.8279 | -105.198
11.8316 | -105.2183
11.8353 | -105.2342
11.8389 | -105.2498
11.8426 | -105.2595
11.8463 | -105.2673
11.8499 | -105.2747
11.8536 | -105.2826
11.8573 | -105.2918
11.861 | -105.2955
11.8646 | -105.2925
11.8683 | -105.2894
11.872 | -105.2845
11.8757 | -105.2761
11.8793 | -105.2692
11.883 | -105.2601
11.8867 | -105.2466
11.8904 | -105.2343
11.894 | -105.2287
11.8977 | -105.2216
11.9014 | -105.2095
11.905 | -105.2021
11.9087 | -105.1972
11.9124 | -105.1941
11.9161 | -105.1901
11.9197 | -105.187
11.9234 | -105.1861
11.9271 | -105.1865
11.9308 | -105.1888
11.9344 | -105.1901
11.9381 | -105.1915
11.9418 | -105.1955
11.9455 | -105.1984
11.9491 | -105.2012
11.9528 | -105.2026
11.9565 | -105.2014
11.9601 | -105.1995
11.9638 | -105.1958
11.9675 | -105.1959
11.9712 | -105.1961
11.9748 | -105.1961
11.9785 | -105.1936
11.9822 | -105.1896
11.9859 | -105.1882
11.9895 | -105.1872
11.9932 | -105.1852
11.9969 | -105.1829
12.0006 | -105.1839
12.0042 | -105.1855
12.0079 | -105.1869
12.0116 | -105.1866
12.0152 | -105.1876
12.0189 | -105.1918
12.0226 | -105.193
12.0263 | -105.1911
12.0299 | -105.1884
12.0336 | -105.1927
12.0373 | -105.1975
12.041 | -105.1997
12.0446 | -105.2047
12.0483 | -105.213
12.052 | -105.2221
12.0556 | -105.2295
12.0593 | -105.2368
12.063 | -105.244
12.0667 | -105.2506
12.0703 | -105.2609
12.074 | -105.2716
12.0777 | -105.2799
12.0814 | -105.2863
12.085 | -105.2947
12.0887 | -105.3043
12.0924 | -105.3129
12.0961 | -105.321
12.0997 | -105.3291
12.1034 | -105.3357
12.1071 | -105.3436
12.1107 | -105.3534
12.1144 | -105.3628
12.1181 | -105.3699
12.1218 | -105.3806
12.1254 | -105.3925
12.1291 | -105.4024
12.1328 | -105.4124
12.1365 | -105.42
12.1401 | -105.428
12.1438 | -105.4364
12.1475 | -105.4436
12.1512 | -105.4528
12.1548 | -105.4613
12.1585 | -105.4634
12.1622 | -105.4656
12.1658 | -105.4709
12.1695 | -105.4755
12.1732 | -105.4779
12.1769 | -105.4824
12.1805 | -105.4834
12.1842 | -105.4858
12.1879 | -105.4874
12.1916 | -105.488
12.1952 | -105.485
12.1989 | -105.4847
12.2026 | -105.4875
12.2063 | -105.4888
12.2099 | -105.4897
12.2136 | -105.4935
12.2173 | -105.4981
12.2209 | -105.5008
12.2246 | -105.5061
12.2283 | -105.5117
12.232 | -105.5196
12.2356 | -105.5311
12.2393 | -105.5411
12.243 | -105.5517
12.2467 | -105.5636
12.2503 | -105.5802
12.254 | -105.5993
12.2577 | -105.6163
12.2614 | -105.6312
12.265 | -105.6474
12.2687 | -105.6618
12.2724 | -105.6765
12.276 | -105.6892
12.2797 | -105.6988
12.2834 | -105.707
12.2871 | -105.7135
12.2907 | -105.7199
12.2944 | -105.725
12.2981 | -105.7282
12.3018 | -105.7297
12.3054 | -105.733
12.3091 | -105.7311
12.3128 | -105.7242
12.3165 | -105.716
12.3201 | -105.7076
12.3238 | -105.6982
12.3275 | -105.6888
12.3311 | -105.6796
12.3348 | -105.6685
12.3385 | -105.6548
12.3422 | -105.6407
12.3458 | -105.6259
12.3495 | -105.6105
12.3532 | -105.5981
12.3569 | -105.5878
12.3605 | -105.5745
12.3642 | -105.563
12.3679 | -105.5537
12.3716 | -105.5465
12.3752 | -105.5417
12.3789 | -105.5396
12.3826 | -105.5364
12.3862 | -105.532
12.3899 | -105.5251
12.3936 | -105.5216
12.3973 | -105.5223
12.4009 | -105.5235
12.4046 | -105.5233
12.4083 | -105.525
12.412 | -105.5271
12.4156 | -105.5262
12.4193 | -105.5281
12.423 | -105.5288
12.4267 | -105.5268
12.4303 | -105.5242
12.434 | -105.5215
12.4377 | -105.5154
12.4413 | -105.5086
12.445 | -105.4997
12.4487 | -105.4908
12.4524 | -105.4794
12.456 | -105.4649
12.4597 | -105.4527
12.4634 | -105.4373
12.4671 | -105.4224
12.4707 | -105.4059
12.4744 | -105.3897
12.4781 | -105.3721
12.4817 | -105.3546
12.4854 | -105.3388
12.4891 | -105.3257
12.4928 | -105.3178
12.4964 | -105.3096
12.5001 | -105.3007
12.5038 | -105.2943
12.5075 | -105.2864
12.5111 | -105.2819
12.5148 | -105.2824
12.5185 | -105.2853
12.5222 | -105.2859
12.5258 | -105.2895
12.5295 | -105.2938
12.5332 | -105.2967
12.5368 | -105.3007
12.5405 | -105.3014
12.5442 | -105.302
12.5479 | -105.3071
12.5515 | -105.3105
12.5552 | -105.3144
12.5589 | -105.3164
12.5626 | -105.3142
12.5662 | -105.3156
12.5699 | -105.3182
12.5736 | -105.3161
12.5773 | -105.3091
12.5809 | -105.3054
12.5846 | -105.3038
12.5883 | -105.3035
12.5919 | -105.3034
12.5956 | -105.3052
12.5993 | -105.3068
12.603 | -105.3099
12.6066 | -105.3162
12.6103 | -105.3208
12.614 | -105.3259
12.6177 | -105.3384
12.6213 | -105.3524
12.625 | -105.368
12.6287 | -105.3862
12.6324 | -105.4029
12.636 | -105.4225
12.6397 | -105.4403
12.6434 | -105.459
12.647 | -105.4769
12.6507 | -105.4957
12.6544 | -105.5149
12.6581 | -105.5315
12.6617 | -105.5494
12.6654 | -105.5665
12.6691 | -105.58
12.6728 | -105.5903
12.6764 | -105.6004
12.6801 | -105.6094
12.6838 | -105.6186
12.6875 | -105.6305
12.6911 | -105.6384
12.6948 | -105.6433
12.6985 | -105.6497
12.7021 | -105.6569
12.7058 | -105.6656
12.7095 | -105.6751
12.7132 | -105.6839
12.7168 | -105.6921
12.7205 | -105.7031
12.7242 | -105.7159
12.7279 | -105.7264
12.7315 | -105.7311
12.7352 | -105.7363
12.7389 | -105.7379
12.7425 | -105.7414
12.7462 | -105.7494
12.7499 | -105.755
12.7536 | -105.759
12.7572 | -105.7638
12.7609 | -105.7667
12.7646 | -105.7645
12.7683 | -105.7559
12.7719 | -105.7491
12.7756 | -105.7417
12.7793 | -105.7309
12.783 | -105.7205
12.7866 | -105.7092
12.7903 | -105.6927
12.794 | -105.6759
12.7976 | -105.6581
12.8013 | -105.6372
12.805 | -105.6144
12.8087 | -105.5916
12.8123 | -105.5672
12.816 | -105.5451
12.8197 | -105.5233
12.8234 | -105.5007
12.827 | -105.4784
12.8307 | -105.4572
12.8344 | -105.437
12.8381 | -105.4176
12.8417 | -105.3999
12.8454 | -105.3799
12.8491 | -105.361
12.8527 | -105.3472
12.8564 | -105.3354
12.8601 | -105.3226
12.8638 | -105.3122
12.8674 | -105.3015
12.8711 | -105.2903
12.8748 | -105.2832
12.8785 | -105.2793
12.8821 | -105.2696
12.8858 | -105.2602
12.8895 | -105.2514
12.8932 | -105.2428
12.8968 | -105.2361
12.9005 | -105.2331
12.9042 | -105.2289
12.9078 | -105.2238
12.9115 | -105.2204
12.9152 | -105.218
12.9189 | -105.2156
12.9225 | -105.2166
12.9262 | -105.2185
12.9299 | -105.2178
12.9336 | -105.2187
12.9372 | -105.2212
12.9409 | -105.2222
12.9446 | -105.2252
12.9483 | -105.2316
12.9519 | -105.2417
12.9556 | -105.2544
12.9593 | -105.2647
12.9629 | -105.2775
12.9666 | -105.2899
12.9703 | -105.3017
12.974 | -105.3177
12.9776 | -105.3374
12.9813 | -105.3551
12.985 | -105.3707
12.9887 | -105.3859
12.9923 | -105.4026
12.996 | -105.4188
12.9997 | -105.4331
13.0034 | -105.4473
13.007 | -105.4589
13.0107 | -105.468
13.0144 | -105.4759
13.018 | -105.4815
13.0217 | -105.4854
13.0254 | -105.489
13.0291 | -105.4911
13.0327 | -105.4912
13.0364 | -105.4877
13.0401 | -105.4791
13.0438 | -105.47
13.0474 | -105.4601
13.0511 | -105.4488
13.0548 | -105.4388
13.0585 | -105.426
13.0621 | -105.4083
13.0658 | -105.3918
13.0695 | -105.3783
13.0731 | -105.3619
13.0768 | -105.3432
13.0805 | -105.3257
13.0842 | -105.3067
13.0878 | -105.2894
13.0915 | -105.2689
13.0952 | -105.2478
13.0989 | -105.2307
13.1025 | -105.2163
13.1062 | -105.2042
13.1099 | -105.1898
13.1136 | -105.178
13.1172 | -105.1685
13.1209 | -105.1584
13.1246 | -105.1508
13.1282 | -105.1482
13.1319 | -105.1482
13.1356 | -105.1493
13.1393 | -105.1503
13.1429 | -105.1561
13.1466 | -105.1626
13.1503 | -105.1685
13.154 | -105.1803
13.1576 | -105.1943
13.1613 | -105.2067
13.165 | -105.2211
13.1686 | -105.2377
13.1723 | -105.2548
13.176 | -105.2733
13.1797 | -105.2908
13.1833 | -105.3063
13.187 | -105.3238
13.1907 | -105.338
13.1944 | -105.3525
13.198 | -105.3665
13.2017 | -105.3778
13.2054 | -105.3928
13.2091 | -105.4064
13.2127 | -105.4123
13.2164 | -105.4173
13.2201 | -105.4228
13.2237 | -105.4276
13.2274 | -105.4319
13.2311 | -105.4316
13.2348 | -105.43
13.2384 | -105.4261
13.2421 | -105.4176
13.2458 | -105.4
13.2495 | -105.3899
13.2531 | -105.3809
13.2568 | -105.3705
13.2605 | -105.3566
13.2642 | -105.3424
13.2678 | -105.3259
13.2715 | -105.3068
13.2752 | -105.2873
13.2788 | -105.2627
13.2825 | -105.2397
13.2862 | -105.2153
13.2899 | -105.1928
13.2935 | -105.1677
13.2972 | -105.144
13.3009 | -105.1183
13.3046 | -105.0906
13.3082 | -105.0663
13.3119 | -105.0471
13.3156 | -105.0244
13.3193 | -105.0008
13.3229 | -104.9788
13.3266 | -104.959
13.3303 | -104.9396
13.3339 | -104.9201
13.3376 | -104.9025
13.3413 | -104.8854
13.345 | -104.8688
13.3486 | -104.8554
13.3523 | -104.8419
13.356 | -104.8285
13.3597 | -104.8209
13.3633 | -104.816
13.367 | -104.8119
13.3707 | -104.8111
13.3744 | -104.8165
13.378 | -104.8241
13.3817 | -104.837
13.3854 | -104.8514
13.389 | -104.8695
13.3927 | -104.8912
13.3964 | -104.9089
13.4001 | -104.9315
13.4037 | -104.9625
13.4074 | -104.9928
13.4111 | -105.0233
13.4148 | -105.0557
13.4184 | -105.0885
13.4221 | -105.1228
13.4258 | -105.1549
13.4294 | -105.1865
13.4331 | -105.2204
13.4368 | -105.2532
13.4405 | -105.2841
13.4441 | -105.3118
13.4478 | -105.3365
13.4515 | -105.3535
13.4552 | -105.368
13.4588 | -105.3849
13.4625 | -105.3977
13.4662 | -105.4047
13.4699 | -105.4081
13.4735 | -105.4085
13.4772 | -105.4057
13.4809 | -105.4022
13.4845 | -105.389
13.4882 | -105.3724
13.4919 | -105.3532
13.4956 | -105.336
13.4992 | -105.3164
13.5029 | -105.291
13.5066 | -105.2612
13.5103 | -105.2339
13.5139 | -105.2057
13.5176 | -105.1732
13.5213 | -105.1402
13.525 | -105.1088
13.5286 | -105.0776
13.5323 | -105.0511
13.536 | -105.0261
13.5396 | -104.9972
13.5433 | -104.9685
13.547 | -104.9443
13.5507 | -104.9235
13.5543 | -104.9021
13.558 | -104.8856
13.5617 | -104.8749
13.5654 | -104.8666
13.569 | -104.8578
13.5727 | -104.851
13.5764 | -104.8449
13.5801 | -104.8423
13.5837 | -104.843
13.5874 | -104.8447
13.5911 | -104.8511
13.5947 | -104.8568
13.5984 | -104.862
13.6021 | -104.8713
13.6058 | -104.8814
13.6094 | -104.8923
13.6131 | -104.9025
13.6168 | -104.9148
13.6205 | -104.9289
13.6241 | -104.9401
13.6278 | -104.9508
13.6315 | -104.9593
13.6352 | -104.9705
13.6388 | -104.9821
13.6425 | -104.9908
13.6462 | -105.0043
13.6498 | -105.0171
13.6535 | -105.0229
13.6572 | -105.0299
13.6609 | -105.038
13.6645 | -105.0441
13.6682 | -105.0472
13.6719 | -105.0497
13.6756 | -105.0525
13.6792 | -105.056
13.6829 | -105.0596
13.6866 | -105.0637
13.6903 | -105.0674
13.6939 | -105.0736
13.6976 | -105.0808
13.7013 | -105.0866
13.7049 | -105.0939
13.7086 | -105.104
13.7123 | -105.1165
13.716 | -105.1288
13.7196 | -105.1403
13.7233 | -105.1525
13.727 | -105.1634
13.7307 | -105.1764
13.7343 | -105.1922
13.738 | -105.2112
13.7417 | -105.23
13.7454 | -105.2461
13.749 | -105.2662
13.7527 | -105.2844
13.7564 | -105.2985
13.76 | -105.3161
13.7637 | -105.3385
13.7674 | -105.3612
13.7711 | -105.3853
13.7747 | -105.4055
13.7784 | -105.4286
13.7821 | -105.4487
13.7858 | -105.4678
13.7894 | -105.4876
13.7931 | -105.5047
13.7968 | -105.5211
13.8004 | -105.5375
13.8041 | -105.5559
13.8078 | -105.5703
13.8115 | -105.584
13.8151 | -105.596
13.8188 | -105.6088
13.8225 | -105.6209
13.8262 | -105.6286
13.8298 | -105.638
13.8335 | -105.6452
13.8372 | -105.6473
13.8409 | -105.6498
13.8445 | -105.6477
13.8482 | -105.6481
13.8519 | -105.6486
13.8555 | -105.6473
13.8592 | -105.6435
13.8629 | -105.6375
13.8666 | -105.6365
13.8702 | -105.6303
13.8739 | -105.6177
13.8776 | -105.6052
13.8813 | -105.5963
13.8849 | -105.5845
13.8886 | -105.5706
13.8923 | -105.5535
13.896 | -105.5362
13.8996 | -105.5221
13.9033 | -105.5085
13.907 | -105.4945
13.9106 | -105.4836
13.9143 | -105.4727
13.918 | -105.4611
13.9217 | -105.4486
13.9253 | -105.4367
13.929 | -105.4253
13.9327 | -105.4133
13.9364 | -105.4001
13.94 | -105.3912
13.9437 | -105.3852
13.9474 | -105.3768
13.9511 | -105.3706
13.9547 | -105.3678
13.9584 | -105.3524
13.9621 | -105.3074
13.9657 | -105.2654
13.9694 | -105.2405
13.9731 | -105.2173
13.9768 | -105.1937
13.9804 | -105.1662
13.9841 | -105.1354
13.9878 | -105.1047
13.9915 | -105.0696
13.9951 | -105.0305
13.9988 | -104.988
14.0025 | -104.9432
14.0062 | -104.8974
14.0098 | -104.8523
14.0135 | -104.8071
14.0172 | -104.7591
14.0208 | -104.709
14.0245 | -104.6622
14.0282 | -104.6169
14.0319 | -104.5679
14.0355 | -104.52
14.0392 | -104.4822
14.0429 | -104.4463
14.0466 | -104.4108
14.0502 | -104.3773
14.0539 | -104.3503
14.0576 | -104.328
14.0613 | -104.3112
14.0649 | -104.3029
14.0686 | -104.2969
14.0723 | -104.2958
14.0759 | -104.3012
14.0796 | -104.3128
14.0833 | -104.3277
14.087 | -104.3482
14.0906 | -104.3746
14.0943 | -104.406
14.098 | -104.4405
14.1017 | -104.4754
14.1053 | -104.5137
14.109 | -104.5522
14.1127 | -104.5892
14.1164 | -104.6303
14.12 | -104.6736
14.1237 | -104.7181
14.1274 | -104.7617
14.131 | -104.8058
14.1347 | -104.852
14.1384 | -104.896
14.1421 | -104.937
14.1457 | -104.9777
14.1494 | -105.0189
14.1531 | -105.0612
14.1568 | -105.1036
14.1604 | -105.1529
14.1641 | -105.2215
14.1678 | -105.2777
14.1715 | -105.3101
14.1751 | -105.3394
14.1788 | -105.3661
14.1825 | -105.3933
14.1861 | -105.4214
14.1898 | -105.4493
14.1935 | -105.4742
14.1972 | -105.5007
14.2008 | -105.5262
14.2045 | -105.5455
14.2082 | -105.5593
14.2119 | -105.5709
14.2155 | -105.5817
14.2192 | -105.5937
14.2229 | -105.6025
14.2265 | -105.6065
14.2302 | -105.6068
14.2339 | -105.6075
14.2376 | -105.6063
14.2412 | -105.5956
14.2449 | -105.5854
14.2486 | -105.5785
14.2523 | -105.5695
14.2559 | -105.5542
14.2596 | -105.5401
14.2633 | -105.5255
14.267 | -105.5074
14.2706 | -105.4901
14.2743 | -105.4744
14.278 | -105.4604
14.2816 | -105.4466
14.2853 | -105.4336
14.289 | -105.4227
14.2927 | -105.411
14.2963 | -105.3982
14.3 | -105.3885
14.3037 | -105.3842
14.3074 | -105.3868
14.311 | -105.3846
14.3147 | -105.3869
14.3184 | -105.3923
14.3221 | -105.3965
14.3257 | -105.4029
14.3294 | -105.4099
14.3331 | -105.4173
14.3367 | -105.4249
14.3404 | -105.4327
14.3441 | -105.4394
14.3478 | -105.4461
14.3514 | -105.4562
14.3551 | -105.4669
14.3588 | -105.4742
14.3625 | -105.478
14.3661 | -105.4837
14.3698 | -105.49
14.3735 | -105.4932
14.3772 | -105.4953
14.3808 | -105.4938
14.3845 | -105.4937
14.3882 | -105.4937
14.3918 | -105.4916
14.3955 | -105.4863
14.3992 | -105.4828
14.4029 | -105.4775
14.4065 | -105.47
14.4102 | -105.4636
14.4139 | -105.4586
14.4176 | -105.4521
14.4212 | -105.4458
14.4249 | -105.4395
14.4286 | -105.435
14.4323 | -105.4305
14.4359 | -105.4299
14.4396 | -105.4327
14.4433 | -105.4311
14.4469 | -105.4298
14.4506 | -105.4296
14.4543 | -105.4288
14.458 | -105.4312
14.4616 | -105.4357
14.4653 | -105.4382
14.469 | -105.4405
14.4727 | -105.4395
14.4763 | -105.4383
14.48 | -105.4422
14.4837 | -105.4436
14.4873 | -105.4408
14.491 | -105.4402
14.4947 | -105.4403
14.4984 | -105.4412
14.502 | -105.4425
14.5057 | -105.4425
14.5094 | -105.4398
14.5131 | -105.4376
14.5167 | -105.4325
14.5204 | -105.4235
14.5241 | -105.4139
14.5278 | -105.4079
14.5314 | -105.4024
14.5351 | -105.3956
14.5388 | -105.3894
14.5424 | -105.3819
14.5461 | -105.3761
14.5498 | -105.3717
14.5535 | -105.3692
14.5571 | -105.3668
14.5608 | -105.3627
14.5645 | -105.3589
14.5682 | -105.3558
14.5718 | -105.3543
14.5755 | -105.3506
14.5792 | -105.346
14.5829 | -105.3461
14.5865 | -105.3464
14.5902 | -105.3487
14.5939 | -105.3483
14.5975 | -105.3482
14.6012 | -105.3477
14.6049 | -105.3509
14.6086 | -105.3536
14.6122 | -105.3575
14.6159 | -105.3611
14.6196 | -105.3671
14.6233 | -105.3738
14.6269 | -105.3826
14.6306 | -105.3902
14.6343 | -105.3971
14.638 | -105.4066
14.6416 | -105.4171
14.6453 | -105.4247
14.649 | -105.4328
14.6526 | -105.4444
14.6563 | -105.4571
14.66 | -105.4724
14.6637 | -105.4876
14.6673 | -105.4978
14.671 | -105.5082
14.6747 | -105.5203
14.6784 | -105.5312
14.682 | -105.5438
14.6857 | -105.5544
14.6894 | -105.5555
14.6931 | -105.5587
14.6967 | -105.5636
14.7004 | -105.5658
14.7041 | -105.5648
14.7077 | -105.5617
14.7114 | -105.5575
14.7151 | -105.5516
14.7188 | -105.5459
14.7224 | -105.5391
14.7261 | -105.5349
14.7298 | -105.5279
14.7335 | -105.5171
14.7371 | -105.5017
14.7408 | -105.4873
14.7445 | -105.4771
14.7482 | -105.4669
14.7518 | -105.4601
14.7555 | -105.4531
14.7592 | -105.4471
14.7628 | -105.4417
14.7665 | -105.4354
14.7702 | -105.4281
14.7739 | -105.4205
14.7775 | -105.4182
14.7812 | -105.417
14.7849 | -105.4164
14.7886 | -105.416
14.7922 | -105.4144
14.7959 | -105.4103
14.7996 | -105.4072
14.8033 | -105.408
14.8069 | -105.4079
14.8106 | -105.4102
14.8143 | -105.4099
14.8179 | -105.4117
14.8216 | -105.4147
14.8253 | -105.4181
14.829 | -105.4176
14.8326 | -105.4154
14.8363 | -105.4144
14.84 | -105.4119
14.8437 | -105.4107
14.8473 | -105.4106
14.851 | -105.4089
14.8547 | -105.4063
14.8583 | -105.4066
14.862 | -105.4022
14.8657 | -105.3982
14.8694 | -105.3957
14.873 | -105.3903
14.8767 | -105.3853
14.8804 | -105.3848
14.8841 | -105.3839
14.8877 | -105.3814
14.8914 | -105.38
14.8951 | -105.3794
14.8988 | -105.3796
14.9024 | -105.3814
14.9061 | -105.3787
14.9098 | -105.3756
14.9134 | -105.374
14.9171 | -105.3716
14.9208 | -105.3717
14.9245 | -105.372
14.9281 | -105.3709
14.9318 | -105.3703
14.9355 | -105.367
14.9392 | -105.3659
14.9428 | -105.365
14.9465 | -105.3652
14.9502 | -105.3666
14.9539 | -105.3649
14.9575 | -105.3612
14.9612 | -105.3577
14.9649 | -105.3554
14.9685 | -105.3516
14.9722 | -105.3471
14.9759 | -105.346
14.9796 | -105.3468
14.9832 | -105.3457
14.9869 | -105.3403
14.9906 | -105.3332
14.9943 | -105.3294
14.9979 | -105.3256
15.0016 | -105.3213
15.0053 | -105.3173
15.009 | -105.3138
15.0126 | -105.3066
15.0163 | -105.299
15.02 | -105.2909
15.0236 | -105.2789
15.0273 | -105.266
15.031 | -105.2541
15.0347 | -105.2415
15.0383 | -105.2307
15.042 | -105.2209
15.0457 | -105.2103
15.0494 | -105.1985
15.053 | -105.1865
15.0567 | -105.1758
15.0604 | -105.1641
15.0641 | -105.1514
15.0677 | -105.1397
15.0714 | -105.1299
15.0751 | -105.1188
15.0787 | -105.1113
15.0824 | -105.1048
15.0861 | -105.0971
15.0898 | -105.0921
15.0934 | -105.0867
15.0971 | -105.0812
15.1008 | -105.0782
15.1045 | -105.0777
15.1081 | -105.0792
15.1118 | -105.081
15.1155 | -105.0827
15.1192 | -105.0866
15.1228 | -105.0911
15.1265 | -105.0971
15.1302 | -105.1055
15.1338 | -105.1161
15.1375 | -105.1255
15.1412 | -105.1361
15.1449 | -105.1467
15.1485 | -105.1558
15.1522 | -105.1667
15.1559 | -105.1716
15.1596 | -105.1779
15.1632 | -105.1877
15.1669 | -105.1967
15.1706 | -105.206
15.1743 | -105.2124
15.1779 | -105.2172
15.1816 | -105.2207
15.1853 | -105.2219
15.1889 | -105.2235
15.1926 | -105.2213
15.1963 | -105.2186
15.2 | -105.22
15.2036 | -105.2237
15.2073 | -105.2247
15.211 | -105.2216
15.2147 | -105.2172
15.2183 | -105.214
15.222 | -105.2097
15.2257 | -105.2092
15.2293 | -105.2091
15.233 | -105.2035
15.2367 | -105.2002
15.2404 | -105.2003
15.244 | -105.2021
15.2477 | -105.2023
15.2514 | -105.202
15.2551 | -105.201
15.2587 | -105.1997
15.2624 | -105.1978
15.2661 | -105.1937
15.2698 | -105.1904
15.2734 | -105.1879
15.2771 | -105.1863
15.2808 | -105.1819
15.2844 | -105.1765
15.2881 | -105.1694
15.2918 | -105.1597
15.2955 | -105.1512
15.2991 | -105.1396
15.3028 | -105.1251
15.3065 | -105.1101
15.3102 | -105.097
15.3138 | -105.0819
15.3175 | -105.0658
15.3212 | -105.0492
15.3249 | -105.0322
15.3285 | -105.0161
15.3322 | -105.0006
15.3359 | -104.9838
15.3395 | -104.9689
15.3432 | -104.9534
15.3469 | -104.9367
15.3506 | -104.9212
15.3542 | -104.9073
15.3579 | -104.8957
15.3616 | -104.8865
15.3653 | -104.8866
15.3689 | -104.8796
15.3726 | -104.8758
15.3763 | -104.8743
15.38 | -104.8745
15.3836 | -104.8721
15.3873 | -104.8723
15.391 | -104.8759
15.3946 | -104.881
15.3983 | -104.8878
15.402 | -104.896
15.4057 | -104.908
15.4093 | -104.9199
15.413 | -104.9309
15.4167 | -104.9424
15.4204 | -104.952
15.424 | -104.9627
15.4277 | -104.9756
15.4314 | -104.9882
15.4351 | -105.0019
15.4387 | -105.014
15.4424 | -105.024
15.4461 | -105.0341
15.4497 | -105.0428
15.4534 | -105.056
15.4571 | -105.0659
15.4608 | -105.0708
15.4644 | -105.0756
15.4681 | -105.0804
15.4718 | -105.0843
15.4755 | -105.0838
15.4791 | -105.0853
15.4828 | -105.0879
15.4865 | -105.089
15.4902 | -105.0877
15.4938 | -105.0842
15.4975 | -105.0784
15.5012 | -105.0724
15.5048 | -105.0653
15.5085 | -105.0568
15.5122 | -105.0469
15.5159 | -105.0361
15.5195 | -105.0273
15.5232 | -105.0173
15.5269 | -105.0053
15.5306 | -104.9905
15.5342 | -104.9742
15.5379 | -104.9603
15.5416 | -104.9465
15.5452 | -104.9333
15.5489 | -104.9185
15.5526 | -104.9039
15.5563 | -104.8915
15.5599 | -104.8805
15.5636 | -104.8713
15.5673 | -104.8608
15.571 | -104.8489
15.5746 | -104.8388
15.5783 | -104.8295
15.582 | -104.8213
15.5857 | -104.8157
15.5893 | -104.8102
15.593 | -104.8067
15.5967 | -104.8049
15.6003 | -104.8049
15.604 | -104.8047
15.6077 | -104.8064
15.6114 | -104.8067
15.615 | -104.8075
15.6187 | -104.8092
15.6224 | -104.8112
15.6261 | -104.8137
15.6297 | -104.8173
15.6334 | -104.8225
15.6371 | -104.8308
15.6408 | -104.8387
15.6444 | -104.8454
15.6481 | -104.8522
15.6518 | -104.8559
15.6554 | -104.865
15.6591 | -104.8749
15.6628 | -104.8843
15.6665 | -104.8956
15.6701 | -104.9031
15.6738 | -104.9094
15.6775 | -104.9155
15.6812 | -104.9228
15.6848 | -104.9332
15.6885 | -104.9381
15.6922 | -104.9422
15.6959 | -104.947
15.6995 | -104.9574
15.7032 | -104.9633
15.7069 | -104.9669
15.7105 | -104.9727
15.7142 | -104.978
15.7179 | -104.9827
15.7216 | -104.9894
15.7252 | -104.9993
15.7289 | -105.0082
15.7326 | -105.0145
15.7363 | -105.0232
15.7399 | -105.0316
15.7436 | -105.0381
15.7473 | -105.0455
15.751 | -105.0535
15.7546 | -105.0655
15.7583 | -105.0755
15.762 | -105.0799
15.7656 | -105.0863
15.7693 | -105.0891
15.773 | -105.0943
15.7767 | -105.1004
15.7803 | -105.1058
15.784 | -105.1102
15.7877 | -105.1124
15.7914 | -105.115
15.795 | -105.1166
15.7987 | -105.1177
15.8024 | -105.1173
15.8061 | -105.1148
15.8097 | -105.1113
15.8134 | -105.1088
15.8171 | -105.1037
15.8207 | -105.0967
15.8244 | -105.0897
15.8281 | -105.084
15.8318 | -105.0788
15.8354 | -105.0732
15.8391 | -105.0691
15.8428 | -105.0654
15.8465 | -105.0619
15.8501 | -105.0587
15.8538 | -105.0545
15.8575 | -105.0502
15.8612 | -105.049
15.8648 | -105.0513
15.8685 | -105.0533
15.8722 | -105.0549
15.8758 | -105.0601
15.8795 | -105.063
15.8832 | -105.0687
15.8869 | -105.0788
15.8905 | -105.0888
15.8942 | -105.103
15.8979 | -105.1278
15.9016 | -105.1461
15.9052 | -105.1629
15.9089 | -105.1827
15.9126 | -105.2015
15.9162 | -105.2212
15.9199 | -105.2433
15.9236 | -105.2668
15.9273 | -105.2862
15.9309 | -105.3061
15.9346 | -105.3257
15.9383 | -105.3454
15.942 | -105.3658
15.9456 | -105.3798
15.9493 | -105.3916
15.953 | -105.4042
15.9567 | -105.4164
15.9603 | -105.4269
15.964 | -105.4349
15.9677 | -105.4375
15.9713 | -105.4418
15.975 | -105.4432
15.9787 | -105.4394
15.9824 | -105.4316
15.986 | -105.4254
15.9897 | -105.42
15.9934 | -105.4081
15.9971 | -105.3958
16.0007 | -105.3811
16.0044 | -105.3646
16.0081 | -105.3467
16.0118 | -105.329
16.0154 | -105.3115
16.0191 | -105.2923
16.0228 | -105.2763
16.0264 | -105.2575
16.0301 | -105.2373
16.0338 | -105.2203
16.0375 | -105.2027
16.0411 | -105.1855
16.0448 | -105.1694
16.0485 | -105.1556
16.0522 | -105.1411
16.0558 | -105.1242
16.0595 | -105.1102
16.0632 | -105.0999
16.0669 | -105.0894
16.0705 | -105.0827
16.0742 | -105.0782
16.0779 | -105.0734
16.0815 | -105.0683
16.0852 | -105.066
16.0889 | -105.0644
16.0926 | -105.0629
16.0962 | -105.0654
16.0999 | -105.0694
16.1036 | -105.0738
16.1073 | -105.0797
16.1109 | -105.0824
16.1146 | -105.0873
16.1183 | -105.0983
16.122 | -105.1086
16.1256 | -105.1161
16.1293 | -105.1223
16.133 | -105.1304
16.1366 | -105.1368
16.1403 | -105.1417
16.144 | -105.1495
16.1477 | -105.1597
16.1513 | -105.1675
16.155 | -105.1718
16.1587 | -105.1777
16.1624 | -105.1828
16.166 | -105.1883
16.1697 | -105.1902
16.1734 | -105.1911
16.1771 | -105.1962
16.1807 | -105.1989
16.1844 | -105.1965
16.1881 | -105.1936
16.1917 | -105.1968
16.1954 | -105.1967
16.1991 | -105.1947
16.2028 | -105.1893
16.2064 | -105.1871
16.2101 | -105.187
16.2138 | -105.186
16.2175 | -105.1846
16.2211 | -105.1794
16.2248 | -105.177
16.2285 | -105.1766
16.2322 | -105.1744
16.2358 | -105.1713
16.2395 | -105.1661
16.2432 | -105.1613
16.2468 | -105.1595
16.2505 | -105.1567
16.2542 | -105.1517
16.2579 | -105.1472
16.2615 | -105.1438
16.2652 | -105.1383
16.2689 | -105.1328
16.2726 | -105.1282
16.2762 | -105.122
16.2799 | -105.1142
16.2836 | -105.1063
16.2872 | -105.1004
16.2909 | -105.095
16.2946 | -105.0902
16.2983 | -105.0859
16.3019 | -105.0829
16.3056 | -105.0807
16.3093 | -105.0796
16.313 | -105.0815
16.3166 | -105.0817
16.3203 | -105.0832
16.324 | -105.0878
16.3277 | -105.0912
16.3313 | -105.0953
16.335 | -105.0996
16.3387 | -105.1065
16.3423 | -105.1174
16.346 | -105.1257
16.3497 | -105.134
16.3534 | -105.1437
16.357 | -105.157
16.3607 | -105.1697
16.3644 | -105.1823
16.3681 | -105.1979
16.3717 | -105.2112
16.3754 | -105.2228
16.3791 | -105.2335
16.3828 | -105.246
16.3864 | -105.2598
16.3901 | -105.2718
16.3938 | -105.2824
16.3974 | -105.292
16.4011 | -105.3029
16.4048 | -105.314
16.4085 | -105.3234
16.4121 | -105.3317
16.4158 | -105.3404
16.4195 | -105.3494
16.4232 | -105.3521
16.4268 | -105.35
16.4305 | -105.3515
16.4342 | -105.3534
16.4379 | -105.3571
16.4415 | -105.3572
16.4452 | -105.355
16.4489 | -105.3511
16.4525 | -105.3454
16.4562 | -105.3397
16.4599 | -105.3327
16.4636 | -105.3265
16.4672 | -105.3192
16.4709 | -105.3117
16.4746 | -105.3051
16.4783 | -105.2954
16.4819 | -105.2838
16.4856 | -105.2746
16.4893 | -105.265
16.493 | -105.2563
16.4966 | -105.2473
16.5003 | -105.2381
16.504 | -105.2288
16.5076 | -105.2192
16.5113 | -105.2116
16.515 | -105.2055
16.5187 | -105.1998
16.5223 | -105.1937
16.526 | -105.1859
16.5297 | -105.1794
16.5334 | -105.1712
16.537 | -105.163
16.5407 | -105.1553
16.5444 | -105.1512
16.5481 | -105.1492
16.5517 | -105.1481
16.5554 | -105.1477
16.5591 | -105.1483
16.5627 | -105.152
16.5664 | -105.1564
16.5701 | -105.1605
16.5738 | -105.1626
16.5774 | -105.1634
16.5811 | -105.1632
16.5848 | -105.1647
16.5885 | -105.1689
16.5921 | -105.1737
16.5958 | -105.1808
16.5995 | -105.1855
16.6031 | -105.1881
16.6068 | -105.1915
16.6105 | -105.1953
16.6142 | -105.2017
16.6178 | -105.2063
16.6215 | -105.2122
16.6252 | -105.2204
16.6289 | -105.2297
16.6325 | -105.2393
16.6362 | -105.2462
16.6399 | -105.2529
16.6436 | -105.2589
16.6472 | -105.2673
16.6509 | -105.2768
16.6546 | -105.2854
16.6582 | -105.2974
16.6619 | -105.3106
16.6656 | -105.3208
16.6693 | -105.3331
16.6729 | -105.345
16.6766 | -105.3544
16.6803 | -105.3613
16.684 | -105.3696
16.6876 | -105.3801
16.6913 | -105.3899
16.695 | -105.3975
16.6987 | -105.4036
16.7023 | -105.4075
16.706 | -105.4132
16.7097 | -105.4191
16.7133 | -105.4245
16.717 | -105.4301
16.7207 | -105.4324
16.7244 | -105.4331
16.728 | -105.4328
16.7317 | -105.4325
16.7354 | -105.4321
16.7391 | -105.4299
16.7427 | -105.4276
16.7464 | -105.4231
16.7501 | -105.4162
16.7538 | -105.4082
16.7574 | -105.4017
16.7611 | -105.3951
16.7648 | -105.3879
16.7684 | -105.3775
16.7721 | -105.3705
16.7758 | -105.3622
16.7795 | -105.354
16.7831 | -105.3436
16.7868 | -105.3305
16.7905 | -105.3192
16.7942 | -105.3079
16.7978 | -105.296
16.8015 | -105.2844
16.8052 | -105.274
16.8089 | -105.2643
16.8125 | -105.2558
16.8162 | -105.249
16.8199 | -105.2454
16.8235 | -105.2372
16.8272 | -105.2278
16.8309 | -105.2192
16.8346 | -105.2113
16.8382 | -105.2075
16.8419 | -105.2043
16.8456 | -105.2016
16.8493 | -105.1985
16.8529 | -105.1944
16.8566 | -105.1917
16.8603 | -105.1902
16.864 | -105.1894
16.8676 | -105.1902
16.8713 | -105.1921
16.875 | -105.1972
16.8786 | -105.2024
16.8823 | -105.2089
16.886 | -105.2151
16.8897 | -105.2203
16.8933 | -105.2278
16.897 | -105.2345
16.9007 | -105.2421
16.9044 | -105.2527
16.908 | -105.2639
16.9117 | -105.2755
16.9154 | -105.2888
16.919 | -105.301
16.9227 | -105.3114
16.9264 | -105.3201
16.9301 | -105.3266
16.9337 | -105.3352
16.9374 | -105.3419
16.9411 | -105.3483
16.9448 | -105.3531
16.9484 | -105.3548
16.9521 | -105.3587
16.9558 | -105.3535
16.9595 | -105.3535
16.9631 | -105.3548
16.9668 | -105.3535
16.9705 | -105.3505
16.9741 | -105.3466
16.9778 | -105.3385
16.9815 | -105.3304
16.9852 | -105.3239
16.9888 | -105.3175
16.9925 | -105.3117
16.9962 | -105.3051
16.9999 | -105.2951
17.0035 | -105.2844
17.0072 | -105.2727
17.0109 | -105.261
17.0146 | -105.2486
17.0182 | -105.2364
17.0219 | -105.2236
17.0256 | -105.2118
17.0292 | -105.2017
17.0329 | -105.192
17.0366 | -105.1805
17.0403 | -105.1693
17.0439 | -105.1588
17.0476 | -105.1482
17.0513 | -105.1386
17.055 | -105.1304
17.0586 | -105.1236
17.0623 | -105.116
17.066 | -105.1125
17.0697 | -105.1131
17.0733 | -105.1133
17.077 | -105.1129
17.0807 | -105.111
17.0843 | -105.1106
17.088 | -105.113
17.0917 | -105.118
17.0954 | -105.1224
17.099 | -105.1267
17.1027 | -105.1324
17.1064 | -105.1386
17.1101 | -105.1431
17.1137 | -105.1498
17.1174 | -105.1563
17.1211 | -105.1618
17.1248 | -105.1684
17.1284 | -105.1734
17.1321 | -105.1777
17.1358 | -105.1795
17.1394 | -105.1833
17.1431 | -105.1854
17.1468 | -105.1865
17.1505 | -105.1855
17.1541 | -105.1839
17.1578 | -105.1793
17.1615 | -105.1778
17.1652 | -105.1751
17.1688 | -105.1696
17.1725 | -105.1648
17.1762 | -105.1565
17.1799 | -105.1462
17.1835 | -105.1382
17.1872 | -105.1312
17.1909 | -105.1243
17.1945 | -105.116
17.1982 | -105.1074
17.2019 | -105.0988
17.2056 | -105.0882
17.2092 | -105.0773
17.2129 | -105.0688
17.2166 | -105.0594
17.2203 | -105.0468
17.2239 | -105.0365
17.2276 | -105.0257
17.2313 | -105.0121
17.235 | -104.9996
17.2386 | -104.986
17.2423 | -104.9707
17.246 | -104.9571
17.2496 | -104.9437
17.2533 | -104.9296
17.257 | -104.918
17.2607 | -104.9065
17.2643 | -104.8936
17.268 | -104.8813
17.2717 | -104.8695
17.2754 | -104.8563
17.279 | -104.8438
17.2827 | -104.8328
17.2864 | -104.8216
17.2901 | -104.8118
17.2937 | -104.803
17.2974 | -104.7936
17.3011 | -104.7844
17.3047 | -104.7747
17.3084 | -104.7647
17.3121 | -104.7537
17.3158 | -104.7424
17.3194 | -104.7322
17.3231 | -104.7246
17.3268 | -104.7183
17.3305 | -104.7125
17.3341 | -104.7043
17.3378 | -104.6963
17.3415 | -104.6884
17.3451 | -104.6817
17.3488 | -104.6743
17.3525 | -104.6639
17.3562 | -104.6572
17.3598 | -104.6499
17.3635 | -104.64
17.3672 | -104.633
17.3709 | -104.6261
17.3745 | -104.6206
17.3782 | -104.6143
17.3819 | -104.6095
17.3856 | -104.6048
17.3892 | -104.6009
17.3929 | -104.5969
17.3966 | -104.5935
17.4002 | -104.5918
17.4039 | -104.5897
17.4076 | -104.5847
17.4113 | -104.5795
17.4149 | -104.5768
17.4186 | -104.5752
17.4223 | -104.5742
17.426 | -104.5741
17.4296 | -104.572
17.4333 | -104.5683
17.437 | -104.5666
17.4407 | -104.5667
17.4443 | -104.5678
17.448 | -104.5706
17.4517 | -104.5735
17.4553 | -104.5768
17.459 | -104.5817
17.4627 | -104.5889
17.4664 | -104.5975
17.47 | -104.6056
17.4737 | -104.6133
17.4774 | -104.6242
17.4811 | -104.6345
17.4847 | -104.6536
17.4884 | -104.6667
17.4921 | -104.682
17.4958 | -104.6979
17.4994 | -104.7129
17.5031 | -104.7283
17.5068 | -104.7427
17.5104 | -104.7581
17.5141 | -104.7752
17.5178 | -104.7913
17.5215 | -104.8061
17.5251 | -104.8224
17.5288 | -104.8381
17.5325 | -104.8566
17.5362 | -104.8743
17.5398 | -104.8871
17.5435 | -104.9011
17.5472 | -104.9153
17.5509 | -104.9262
17.5545 | -104.9373
17.5582 | -104.947
17.5619 | -104.9523
17.5655 | -104.9567
17.5692 | -104.9592
17.5729 | -104.9565
17.5766 | -104.9544
17.5802 | -104.9524
17.5839 | -104.9475
17.5876 | -104.9396
17.5913 | -104.9316
17.5949 | -104.9211
17.5986 | -104.9064
17.6023 | -104.8933
17.606 | -104.8763
17.6096 | -104.8558
17.6133 | -104.8363
17.617 | -104.8129
17.6206 | -104.7873
17.6243 | -104.7608
17.628 | -104.7374
17.6317 | -104.712
17.6353 | -104.6839
17.639 | -104.657
17.6427 | -104.6322
17.6464 | -104.609
17.65 | -104.5843
17.6537 | -104.561
17.6574 | -104.5384
17.661 | -104.5184
17.6647 | -104.4987
17.6684 | -104.4789
17.6721 | -104.4606
17.6757 | -104.4442
17.6794 | -104.4282
17.6831 | -104.4162
17.6868 | -104.405
17.6904 | -104.3915
17.6941 | -104.3796
17.6978 | -104.3715
17.7015 | -104.3655
17.7051 | -104.3603
17.7088 | -104.3555
17.7125 | -104.3511
17.7161 | -104.3495
17.7198 | -104.3506
17.7235 | -104.3529
17.7272 | -104.3539
17.7308 | -104.3545
17.7345 | -104.3547
17.7382 | -104.3544
17.7419 | -104.3544
17.7455 | -104.3535
17.7492 | -104.3542
17.7529 | -104.3562
17.7566 | -104.3584
17.7602 | -104.3607
17.7639 | -104.3605
17.7676 | -104.3587
17.7712 | -104.3579
17.7749 | -104.3557
17.7786 | -104.3527
17.7823 | -104.3496
17.7859 | -104.3443
17.7896 | -104.3385
17.7933 | -104.3359
17.797 | -104.3301
17.8006 | -104.3225
17.8043 | -104.3166
17.808 | -104.3096
17.8117 | -104.3023
17.8153 | -104.2966
17.819 | -104.2891
17.8227 | -104.2802
17.8263 | -104.2747
17.83 | -104.2669
17.8337 | -104.2568
17.8374 | -104.2496
17.841 | -104.2434
17.8447 | -104.2347
17.8484 | -104.2286
17.8521 | -104.2234
17.8557 | -104.2186
17.8594 | -104.2135
17.8631 | -104.2105
17.8668 | -104.2067
17.8704 | -104.2015
17.8741 | -104.1979
17.8778 | -104.1914
17.8814 | -104.1812
17.8851 | -104.1736
17.8888 | -104.1655
17.8925 | -104.1566
17.8961 | -104.1469
17.8998 | -104.1379
17.9035 | -104.1274
17.9072 | -104.1136
17.9108 | -104.1005
17.9145 | -104.0817
17.9182 | -104.0581
17.9219 | -104.037
17.9255 | -104.016
17.9292 | -103.994
17.9329 | -103.9688
17.9365 | -103.9423
17.9402 | -103.9148
17.9439 | -103.8885
17.9476 | -103.8624
17.9512 | -103.8325
17.9549 | -103.8001
17.9586 | -103.7682
17.9623 | -103.7357
17.9659 | -103.7044
17.9696 | -103.6729
17.9733 | -103.6402
17.9769 | -103.6079
17.9806 | -103.5763
17.9843 | -103.5462
17.988 | -103.5181
17.9916 | -103.4896
17.9953 | -103.4591
17.999 | -103.4295
18.0027 | -103.4027
18.0063 | -103.3782
18.01 | -103.3562
18.0137 | -103.3353
18.0174 | -103.3268
18.021 | -103.3079
18.0247 | -103.2921
18.0284 | -103.2776
18.032 | -103.2663
18.0357 | -103.2578
18.0394 | -103.251
18.0431 | -103.246
18.0467 | -103.2439
18.0504 | -103.2435
18.0541 | -103.2447
18.0578 | -103.2495
18.0614 | -103.2564
18.0651 | -103.2641
18.0688 | -103.2735
18.0725 | -103.2844
18.0761 | -103.2976
18.0798 | -103.3137
18.0835 | -103.3306
18.0871 | -103.3488
18.0908 | -103.3664
18.0945 | -103.3852
18.0982 | -103.4057
18.1018 | -103.4285
18.1055 | -103.4496
18.1092 | -103.471
18.1129 | -103.4924
18.1165 | -103.5159
18.1202 | -103.5398
18.1239 | -103.5646
18.1276 | -103.5905
18.1312 | -103.6131
18.1349 | -103.6396
18.1386 | -103.6674
18.1422 | -103.6924
18.1459 | -103.7166
18.1496 | -103.7383
18.1533 | -103.7617
18.1569 | -103.7836
18.1606 | -103.8065
18.1643 | -103.8312
18.168 | -103.8545
18.1716 | -103.8792
18.1753 | -103.9028
18.179 | -103.9217
18.1827 | -103.9409
18.1863 | -103.9562
18.19 | -103.9719
18.1937 | -103.9859
18.1973 | -103.9996
18.201 | -104.0128
18.2047 | -104.0241
18.2084 | -104.0351
18.212 | -104.046
18.2157 | -104.0538
18.2194 | -104.0602
18.2231 | -104.0649
18.2267 | -104.0668
18.2304 | -104.0705
18.2341 | -104.0742
18.2378 | -104.076
18.2414 | -104.0786
18.2451 | -104.0801
18.2488 | -104.0832
18.2524 | -104.0883
18.2561 | -104.0895
18.2598 | -104.0856
18.2635 | -104.0819
18.2671 | -104.0778
18.2708 | -104.0736
18.2745 | -104.0708
18.2782 | -104.0649
18.2818 | -104.0591
18.2855 | -104.0507
18.2892 | -104.04
18.2929 | -104.0291
18.2965 | -104.0173
18.3002 | -104.0038
18.3039 | -103.9878
18.3075 | -103.9726
18.3112 | -103.9575
18.3149 | -103.94
18.3186 | -103.9207
18.3222 | -103.8997
18.3259 | -103.883
18.3296 | -103.8663
18.3333 | -103.8489
18.3369 | -103.8307
18.3406 | -103.8125
18.3443 | -103.795
18.348 | -103.7784
18.3516 | -103.7658
18.3553 | -103.7527
18.359 | -103.7359
18.3626 | -103.7215
18.3663 | -103.7093
18.37 | -103.6996
18.3737 | -103.6906
18.3773 | -103.6822
18.381 | -103.6751
18.3847 | -103.6693
18.3884 | -103.6644
18.392 | -103.6638
18.3957 | -103.6623
18.3994 | -103.6624
18.403 | -103.6639
18.4067 | -103.6671
18.4104 | -103.6719
18.4141 | -103.6748
18.4177 | -103.6797
18.4214 | -103.6848
18.4251 | -103.691
18.4288 | -103.6979
18.4324 | -103.7044
18.4361 | -103.7131
18.4398 | -103.7228
18.4435 | -103.7314
18.4471 | -103.7413
18.4508 | -103.7508
18.4545 | -103.7596
18.4581 | -103.7687
18.4618 | -103.7752
18.4655 | -103.7828
18.4692 | -103.7894
18.4728 | -103.7967
18.4765 | -103.8044
18.4802 | -103.8117
18.4839 | -103.8172
18.4875 | -103.8231
18.4912 | -103.827
18.4949 | -103.8273
18.4986 | -103.8278
18.5022 | -103.8258
18.5059 | -103.8237
18.5096 | -103.8209
18.5132 | -103.8173
18.5169 | -103.8136
18.5206 | -103.8088
18.5243 | -103.8004
18.5279 | -103.7892
18.5316 | -103.7788
18.5353 | -103.7671
18.539 | -103.7565
18.5426 | -103.7454
18.5463 | -103.7275
18.55 | -103.7154
18.5537 | -103.7022
18.5573 | -103.6891
18.561 | -103.6746
18.5647 | -103.6579
18.5683 | -103.6432
18.572 | -103.6276
18.5757 | -103.6107
18.5794 | -103.5951
18.583 | -103.5809
18.5867 | -103.567
18.5904 | -103.5484
18.5941 | -103.5288
18.5977 | -103.5119
18.6014 | -103.497
18.6051 | -103.483
18.6088 | -103.468
18.6124 | -103.4541
18.6161 | -103.4406
18.6198 | -103.4283
18.6234 | -103.4188
18.6271 | -103.4095
18.6308 | -103.4004
18.6345 | -103.3917
18.6381 | -103.3817
18.6418 | -103.3712
18.6455 | -103.3616
18.6492 | -103.3528
18.6528 | -103.3448
18.6565 | -103.3384
18.6602 | -103.3342
18.6639 | -103.3279
18.6675 | -103.3233
18.6712 | -103.3207
18.6749 | -103.3185
18.6785 | -103.3191
18.6822 | -103.3199
18.6859 | -103.321
18.6896 | -103.3218
18.6932 | -103.3234
18.6969 | -103.3264
18.7006 | -103.3307
18.7043 | -103.335
18.7079 | -103.3368
18.7116 | -103.3382
18.7153 | -103.3424
18.7189 | -103.3475
18.7226 | -103.3534
18.7263 | -103.3622
18.73 | -103.3666
18.7336 | -103.3724
18.7373 | -103.3774
18.741 | -103.3828
18.7447 | -103.3867
18.7483 | -103.3898
18.752 | -103.394
18.7557 | -103.3976
18.7594 | -103.4022
18.763 | -103.4067
18.7667 | -103.4105
18.7704 | -103.4105
18.774 | -103.4084
18.7777 | -103.4067
18.7814 | -103.4064
18.7851 | -103.4064
18.7887 | -103.4051
18.7924 | -103.4049
18.7961 | -103.4023
18.7998 | -103.3995
18.8034 | -103.3963
18.8071 | -103.3926
18.8108 | -103.387
18.8145 | -103.3816
18.8181 | -103.3775
18.8218 | -103.373
18.8255 | -103.3668
18.8291 | -103.362
18.8328 | -103.3588
18.8365 | -103.3546
18.8402 | -103.3506
18.8438 | -103.3478
18.8475 | -103.3458
18.8512 | -103.3435
18.8549 | -103.3413
18.8585 | -103.3409
18.8622 | -103.3424
18.8659 | -103.3423
18.8696 | -103.3408
18.8732 | -103.3412
18.8769 | -103.3424
18.8806 | -103.3429
18.8842 | -103.3429
18.8879 | -103.3435
18.8916 | -103.3476
18.8953 | -103.3504
18.8989 | -103.3549
18.9026 | -103.3593
18.9063 | -103.3649
18.91 | -103.3699
18.9136 | -103.3741
18.9173 | -103.3793
18.921 | -103.3854
18.9247 | -103.3895
18.9283 | -103.3932
18.932 | -103.3978
18.9357 | -103.4013
18.9393 | -103.4034
18.943 | -103.4047
18.9467 | -103.4046
18.9504 | -103.4042
18.954 | -103.4033
18.9577 | -103.4009
18.9614 | -103.4005
18.9651 | -103.3999
18.9687 | -103.3987
18.9724 | -103.3966
18.9761 | -103.3969
18.9798 | -103.3996
18.9834 | -103.3981
18.9871 | -103.3942
18.9908 | -103.3885
18.9944 | -103.3832
18.9981 | -103.3815
19.0018 | -103.3769
19.0055 | -103.3699
19.0091 | -103.3638
19.0128 | -103.3587
19.0165 | -103.3549
19.0202 | -103.3488
19.0238 | -103.3402
19.0275 | -103.3321
19.0312 | -103.3226
19.0348 | -103.3119
19.0385 | -103.3025
19.0422 | -103.2956
19.0459 | -103.2879
19.0495 | -103.2773
19.0532 | -103.2648
19.0569 | -103.2523
19.0606 | -103.2388
19.0642 | -103.2262
19.0679 | -103.2142
19.0716 | -103.2008
19.0753 | -103.1867
19.0789 | -103.1719
19.0826 | -103.1581
19.0863 | -103.1456
19.0899 | -103.1312
19.0936 | -103.1181
19.0973 | -103.1051
19.101 | -103.0903
19.1046 | -103.0768
19.1083 | -103.0661
19.112 | -103.0561
19.1157 | -103.046
19.1193 | -103.0387
19.123 | -103.0318
19.1267 | -103.0247
19.1304 | -103.0192
19.134 | -103.0146
19.1377 | -103.0115
19.1414 | -103.0097
19.145 | -103.0087
19.1487 | -103.009
19.1524 | -103.0121
19.1561 | -103.0164
19.1597 | -103.02
19.1634 | -103.0244
19.1671 | -103.0291
19.1708 | -103.0339
19.1744 | -103.0397
19.1781 | -103.0473
19.1818 | -103.0547
19.1855 | -103.0618
19.1891 | -103.0684
19.1928 | -103.0733
19.1965 | -103.0801
19.2001 | -103.0873
19.2038 | -103.0927
19.2075 | -103.0982
19.2112 | -103.1062
19.2148 | -103.1157
19.2185 | -103.1225
19.2222 | -103.1293
19.2259 | -103.1364
19.2295 | -103.1441
19.2332 | -103.1489
19.2369 | -103.1524
19.2406 | -103.154
19.2442 | -103.1546
19.2479 | -103.1565
19.2516 | -103.1568
19.2552 | -103.1586
19.2589 | -103.1601
19.2626 | -103.1601
19.2663 | -103.1602
19.2699 | -103.1613
19.2736 | -103.1618
19.2773 | -103.1623
19.281 | -103.1621
19.2846 | -103.1636
19.2883 | -103.1628
19.292 | -103.1635
19.2957 | -103.1671
19.2993 | -103.1715
19.303 | -103.176
19.3067 | -103.1807
19.3103 | -103.1859
19.314 | -103.1899
19.3177 | -103.1933
19.3214 | -103.1979
19.325 | -103.2046
19.3287 | -103.2145
19.3324 | -103.2261
19.3361 | -103.237
19.3397 | -103.2467
19.3434 | -103.2579
19.3471 | -103.2684
19.3508 | -103.2796
19.3544 | -103.2933
19.3581 | -103.3075
19.3618 | -103.3194
19.3654 | -103.331
19.3691 | -103.3428
19.3728 | -103.3535
19.3765 | -103.3663
19.3801 | -103.3778
19.3838 | -103.387
19.3875 | -103.3949
19.3912 | -103.4012
19.3948 | -103.4075
19.3985 | -103.4133
19.4022 | -103.4184
19.4059 | -103.4214
19.4095 | -103.4255
19.4132 | -103.4293
19.4169 | -103.4322
19.4205 | -103.4339
19.4242 | -103.4342
19.4279 | -103.4323
19.4316 | -103.4309
19.4352 | -103.4301
19.4389 | -103.4257
19.4426 | -103.4227
19.4463 | -103.4184
19.4499 | -103.413
19.4536 | -103.4074
19.4573 | -103.4021
19.4609 | -103.3967
19.4646 | -103.3926
19.4683 | -103.3886
19.472 | -103.3802
19.4756 | -103.3708
19.4793 | -103.3649
19.483 | -103.3602
19.4867 | -103.353
19.4903 | -103.3439
19.494 | -103.3378
19.4977 | -103.3317
19.5014 | -103.3272
19.505 | -103.3235
19.5087 | -103.3211
19.5124 | -103.3203
19.516 | -103.3212
19.5197 | -103.3243
19.5234 | -103.3258
19.5271 | -103.3281
19.5307 | -103.3314
19.5344 | -103.3364
19.5381 | -103.3421
19.5418 | -103.349
19.5454 | -103.353
19.5491 | -103.3603
19.5528 | -103.3709
19.5565 | -103.3809
19.5601 | -103.3895
19.5638 | -103.401
19.5675 | -103.4122
19.5711 | -103.4217
19.5748 | -103.4326
19.5785 | -103.446
19.5822 | -103.4586
19.5858 | -103.4677
19.5895 | -103.4767
19.5932 | -103.4854
19.5969 | -103.4967
19.6005 | -103.5068
19.6042 | -103.5137
19.6079 | -103.5201
19.6116 | -103.5233
19.6152 | -103.528
19.6189 | -103.533
19.6226 | -103.5368
19.6262 | -103.5388
19.6299 | -103.5397
19.6336 | -103.5392
19.6373 | -103.5373
19.6409 | -103.5327
19.6446 | -103.5287
19.6483 | -103.5203
19.652 | -103.5122
19.6556 | -103.5046
19.6593 | -103.4944
19.663 | -103.4821
19.6667 | -103.4694
19.6703 | -103.4576
19.674 | -103.4412
19.6777 | -103.4232
19.6813 | -103.4055
19.685 | -103.3888
19.6887 | -103.3702
19.6924 | -103.3501
19.696 | -103.3293
19.6997 | -103.3097
19.7034 | -103.2906
19.7071 | -103.2709
19.7107 | -103.2522
19.7144 | -103.235
19.7181 | -103.2179
19.7218 | -103.199
19.7254 | -103.1818
19.7291 | -103.1672
19.7328 | -103.1523
19.7364 | -103.1374
19.7401 | -103.1236
19.7438 | -103.1102
19.7475 | -103.1002
19.7511 | -103.0914
19.7548 | -103.0813
19.7585 | -103.0735
19.7622 | -103.0664
19.7658 | -103.0585
19.7695 | -103.053
19.7732 | -103.0483
19.7768 | -103.0452
19.7805 | -103.0423
19.7842 | -103.0413
19.7879 | -103.0406
19.7915 | -103.0418
19.7952 | -103.0437
19.7989 | -103.0467
19.8026 | -103.0515
19.8062 | -103.0577
19.8099 | -103.0632
19.8136 | -103.0682
19.8173 | -103.0759
19.8209 | -103.0822
19.8246 | -103.0906
19.8283 | -103.1004
19.8319 | -103.1082
19.8356 | -103.1146
19.8393 | -103.1216
19.843 | -103.1282
19.8466 | -103.1365
19.8503 | -103.1446
19.854 | -103.1507
19.8577 | -103.1571
19.8613 | -103.1664
19.865 | -103.1775
19.8687 | -103.186
19.8724 | -103.1951
19.876 | -103.2016
19.8797 | -103.2069
19.8834 | -103.2111
19.887 | -103.2141
19.8907 | -103.221
19.8944 | -103.2297
19.8981 | -103.2361
19.9017 | -103.2424
19.9054 | -103.2502
19.9091 | -103.2588
19.9128 | -103.2654
19.9164 | -103.2706
19.9201 | -103.2765
19.9238 | -103.284
19.9275 | -103.2905
19.9311 | -103.2985
19.9348 | -103.3046
19.9385 | -103.3063
19.9421 | -103.3115
19.9458 | -103.3179
19.9495 | -103.3242
19.9532 | -103.3284
19.9568 | -103.3325
19.9605 | -103.3369
19.9642 | -103.3396
19.9679 | -103.3398
19.9715 | -103.3395
19.9752 | -103.3396
19.9789 | -103.3399
19.9826 | -103.3383
19.9862 | -103.3364
19.9899 | -103.3347
19.9936 | -103.3313
19.9972 | -103.3273
20.0009 | -103.3244
20.0046 | -103.3229
20.0083 | -103.3201
20.0119 | -103.3161
20.0156 | -103.312
20.0193 | -103.3088
20.023 | -103.3032
20.0266 | -103.2968
20.0303 | -103.2886
20.034 | -103.2812
20.0377 | -103.276
20.0413 | -103.2697
20.045 | -103.2626
20.0487 | -103.2555
20.0523 | -103.2473
20.056 | -103.2378
20.0597 | -103.2288
20.0634 | -103.2198
20.067 | -103.2111
20.0707 | -103.2024
20.0744 | -103.194
20.0781 | -103.1867
20.0817 | -103.178
20.0854 | -103.1699
20.0891 | -103.1621
20.0927 | -103.1554
20.0964 | -103.15
20.1001 | -103.1423
20.1038 | -103.1347
20.1074 | -103.1274
20.1111 | -103.1188
20.1148 | -103.1125
20.1185 | -103.1091
20.1221 | -103.1049
20.1258 | -103.1009
20.1295 | -103.0971
20.1332 | -103.095
20.1368 | -103.0944
20.1405 | -103.0908
20.1442 | -103.0865
20.1478 | -103.082
20.1515 | -103.0791
20.1552 | -103.0792
20.1589 | -103.0774
20.1625 | -103.0753
20.1662 | -103.075
20.1699 | -103.0748
20.1736 | -103.0749
20.1772 | -103.0728
20.1809 | -103.0709
20.1846 | -103.0694
20.1883 | -103.0687
20.1919 | -103.0665
20.1956 | -103.0649
20.1993 | -103.0665
20.2029 | -103.068
20.2066 | -103.0686
20.2103 | -103.0682
20.214 | -103.0649
20.2176 | -103.0635
20.2213 | -103.0617
20.225 | -103.0593
20.2287 | -103.055
20.2323 | -103.0494
20.236 | -103.0423
20.2397 | -103.0367
20.2434 | -103.0298
20.247 | -103.0232
20.2507 | -103.018
20.2544 | -103.0153
20.258 | -103.0117
20.2617 | -103.0056
20.2654 | -102.9987
20.2691 | -102.9905
20.2727 | -102.9837
20.2764 | -102.9779
20.2801 | -102.9715
20.2838 | -102.9645
20.2874 | -102.9562
20.2911 | -102.9484
20.2948 | -102.9395
20.2985 | -102.929
20.3021 | -102.9192
20.3058 | -102.9109
20.3095 | -102.9039
20.3131 | -102.8961
20.3168 | -102.8873
20.3205 | -102.8783
20.3242 | -102.8704
20.3278 | -102.8626
20.3315 | -102.8536
20.3352 | -102.8465
20.3389 | -102.8387
20.3425 | -102.8338
20.3462 | -102.829
20.3499 | -102.824
20.3536 | -102.8182
20.3572 | -102.8131
20.3609 | -102.8082
20.3646 | -102.8066
20.3682 | -102.8051
20.3719 | -102.8001
20.3756 | -102.7963
20.3793 | -102.7954
20.3829 | -102.7975
20.3866 | -102.7999
20.3903 | -102.8022
20.394 | -102.8064
20.3976 | -102.8115
20.4013 | -102.8153
20.405 | -102.8191
20.4087 | -102.8232
20.4123 | -102.8293
20.416 | -102.8344
20.4197 | -102.8397
20.4233 | -102.8461
20.427 | -102.8523
20.4307 | -102.8567
20.4344 | -102.8618
20.438 | -102.8678
20.4417 | -102.8735
20.4454 | -102.8764
20.4491 | -102.8817
20.4527 | -102.8867
20.4564 | -102.8883
20.4601 | -102.8889
20.4637 | -102.8886
20.4674 | -102.8885
20.4711 | -102.8881
20.4748 | -102.889
20.4784 | -102.8875
20.4821 | -102.887
20.4858 | -102.8839
20.4895 | -102.8783
20.4931 | -102.8737
20.4968 | -102.8683
20.5005 | -102.8623
20.5042 | -102.8589
20.5078 | -102.8549
20.5115 | -102.8514
20.5152 | -102.8483
20.5188 | -102.8428
20.5225 | -102.8364
20.5262 | -102.8308
20.5299 | -102.8256
20.5335 | -102.8211
20.5372 | -102.8144
20.5409 | -102.8067
20.5446 | -102.7989
20.5482 | -102.7921
20.5519 | -102.7865
20.5556 | -102.7821
20.5593 | -102.7779
20.5629 | -102.7719
20.5666 | -102.7664
20.5703 | -102.7616
20.5739 | -102.7581
20.5776 | -102.7574
20.5813 | -102.7574
20.585 | -102.7598
20.5886 | -102.7634
20.5923 | -102.7669
20.596 | -102.7675
20.5997 | -102.768
20.6033 | -102.7699
20.607 | -102.7742
20.6107 | -102.7778
20.6144 | -102.7798
20.618 | -102.779
20.6217 | -102.7776
20.6254 | -102.7763
20.629 | -102.7759
20.6327 | -102.7717
20.6364 | -102.7685
20.6401 | -102.7652
20.6437 | -102.7573
20.6474 | -102.749
20.6511 | -102.7426
20.6548 | -102.7347
20.6584 | -102.726
20.6621 | -102.7181
20.6658 | -102.6975
20.6695 | -102.6876
20.6731 | -102.6749
20.6768 | -102.6605
20.6805 | -102.6465
20.6841 | -102.6328
20.6878 | -102.6173
20.6915 | -102.5995
20.6952 | -102.5811
20.6988 | -102.5614
20.7025 | -102.5428
20.7062 | -102.5224
20.7099 | -102.5
20.7135 | -102.4773
20.7172 | -102.4545
20.7209 | -102.4339
20.7246 | -102.4162
20.7282 | -102.3977
20.7319 | -102.3772
20.7356 | -102.3578
20.7392 | -102.3387
20.7429 | -102.32
20.7466 | -102.3035
20.7503 | -102.2875
20.7539 | -102.2721
20.7576 | -102.2583
20.7613 | -102.2456
20.765 | -102.2346
20.7686 | -102.224
20.7723 | -102.2158
20.776 | -102.2081
20.7797 | -102.2009
20.7833 | -102.1953
20.787 | -102.1898
20.7907 | -102.1857
20.7943 | -102.1843
20.798 | -102.1828
20.8017 | -102.1796
20.8054 | -102.1783
20.809 | -102.1761
20.8127 | -102.1757
20.8164 | -102.179
20.8201 | -102.1815
20.8237 | -102.1821
20.8274 | -102.1837
20.8311 | -102.184
20.8347 | -102.1846
20.8384 | -102.1835
20.8421 | -102.1824
20.8458 | -102.1818
20.8494 | -102.1798
20.8531 | -102.1798
20.8568 | -102.1796
20.8605 | -102.1774
20.8641 | -102.174
20.8678 | -102.1699
20.8715 | -102.1635
20.8752 | -102.1557
20.8788 | -102.1472
20.8825 | -102.139
20.8862 | -102.1309
20.8898 | -102.1217
20.8935 | -102.1138
20.8972 | -102.1049
20.9009 | -102.0953
20.9045 | -102.0827
20.9082 | -102.0708
20.9119 | -102.0615
20.9156 | -102.054
20.9192 | -102.0465
20.9229 | -102.0386
20.9266 | -102.0288
20.9303 | -102.0185
20.9339 | -102.0072
20.9376 | -101.9977
20.9413 | -101.9894
20.9449 | -101.9782
20.9486 | -101.9674
20.9523 | -101.9595
20.956 | -101.9517
20.9596 | -101.9439
20.9633 | -101.9371
20.967 | -101.9309
20.9707 | -101.9246
20.9743 | -101.917
20.978 | -101.9124
20.9817 | -101.9101
20.9854 | -101.9074
20.989 | -101.9066
20.9927 | -101.9064
20.9964 | -101.9074
21 | -101.9083
21.0037 | -101.9109
21.0074 | -101.9152
21.0111 | -101.9202
21.0147 | -101.9238
21.0184 | -101.9266
21.0221 | -101.9299
21.0258 | -101.9323
21.0294 | -101.9361
21.0331 | -101.9396
21.0368 | -101.9454
21.0405 | -101.9504
21.0441 | -101.9542
21.0478 | -101.9586
21.0515 | -101.9637
21.0551 | -101.9676
21.0588 | -101.9723
21.0625 | -101.9773
21.0662 | -101.982
21.0698 | -101.9868
21.0735 | -101.9913
21.0772 | -101.998
21.0809 | -102.0056
21.0845 | -102.0128
21.0882 | -102.0184
21.0919 | -102.022
21.0956 | -102.0268
21.0992 | -102.0334
21.1029 | -102.0384
21.1066 | -102.045
21.1102 | -102.0502
21.1139 | -102.052
21.1176 | -102.0544
21.1213 | -102.0554
21.1249 | -102.0561
21.1286 | -102.0593
21.1323 | -102.062
21.136 | -102.0632
21.1396 | -102.0643
21.1433 | -102.0647
21.147 | -102.0643
21.1506 | -102.0618
21.1543 | -102.0599
21.158 | -102.0591
21.1617 | -102.0585
21.1653 | -102.0558
21.169 | -102.0533
21.1727 | -102.0506
21.1764 | -102.0463
21.18 | -102.0433
21.1837 | -102.04
21.1874 | -102.038
21.1911 | -102.0335
21.1947 | -102.0294
21.1984 | -102.0277
21.2021 | -102.0253
21.2057 | -102.0258
21.2094 | -102.0252
21.2131 | -102.0236
21.2168 | -102.022
21.2204 | -102.0203
21.2241 | -102.0194
21.2278 | -102.0212
21.2315 | -102.0215
21.2351 | -102.0234
21.2388 | -102.0259
21.2425 | -102.0266
21.2462 | -102.0295
21.2498 | -102.0315
21.2535 | -102.0342
21.2572 | -102.0391
21.2608 | -102.0441
21.2645 | -102.0468
21.2682 | -102.0501
21.2719 | -102.0555
21.2755 | -102.0634
21.2792 | -102.0696
21.2829 | -102.0743
21.2866 | -102.0809
21.2902 | -102.087
21.2939 | -102.0936
21.2976 | -102.0998
21.3013 | -102.1062
21.3049 | -102.1115
21.3086 | -102.1173
21.3123 | -102.1219
21.3159 | -102.1283
21.3196 | -102.1339
21.3233 | -102.1387
21.327 | -102.1437
21.3306 | -102.1477
21.3343 | -102.1513
21.338 | -102.1556
21.3417 | -102.1603
21.3453 | -102.1645
21.349 | -102.1689
21.3527 | -102.1723
21.3564 | -102.1754
21.36 | -102.1784
21.3637 | -102.1797
21.3674 | -102.1829
21.371 | -102.1869
21.3747 | -102.1916
21.3784 | -102.1969
21.3821 | -102.2008
21.3857 | -102.2044
21.3894 | -102.2111
21.3931 | -102.2184
21.3968 | -102.2244
21.4004 | -102.2308
21.4041 | -102.2395
21.4078 | -102.2486
21.4115 | -102.2595
21.4151 | -102.2699
21.4188 | -102.2807
21.4225 | -102.2945
21.4261 | -102.3093
21.4298 | -102.3229
21.4335 | -102.3353
21.4372 | -102.3488
21.4408 | -102.3607
21.4445 | -102.3748
21.4482 | -102.3906
21.4519 | -102.4059
21.4555 | -102.423
21.4592 | -102.4394
21.4629 | -102.4541
21.4666 | -102.4715
21.4702 | -102.4897
21.4739 | -102.5083
21.4776 | -102.5279
21.4812 | -102.5435
21.4849 | -102.5581
21.4886 | -102.5717
21.4923 | -102.5858
21.4959 | -102.6006
21.4996 | -102.6134
21.5033 | -102.6229
21.507 | -102.6338
21.5106 | -102.6429
21.5143 | -102.6522
21.518 | -102.661
21.5216 | -102.6692
21.5253 | -102.6745
21.529 | -102.679
21.5327 | -102.682
21.5363 | -102.6869
21.54 | -102.6892
21.5437 | -102.692
21.5474 | -102.6949
21.551 | -102.6952
21.5547 | -102.6948
21.5584 | -102.6923
21.5621 | -102.6882
21.5657 | -102.6841
21.5694 | -102.6789
21.5731 | -102.6713
21.5767 | -102.665
21.5804 | -102.6567
21.5841 | -102.6472
21.5878 | -102.6378
21.5914 | -102.6295
21.5951 | -102.622
21.5988 | -102.6152
21.6025 | -102.6085
21.6061 | -102.6011
21.6098 | -102.5941
21.6135 | -102.5872
21.6172 | -102.5808
21.6208 | -102.5745
21.6245 | -102.5676
21.6282 | -102.5626
21.6318 | -102.5591
21.6355 | -102.5565
21.6392 | -102.5552
21.6429 | -102.5539
21.6465 | -102.5537
21.6502 | -102.556
21.6539 | -102.5584
21.6576 | -102.561
21.6612 | -102.5649
21.6649 | -102.5685
21.6686 | -102.5735
21.6723 | -102.5797
21.6759 | -102.5872
21.6796 | -102.5968
21.6833 | -102.607
21.6869 | -102.6189
21.6906 | -102.6333
21.6943 | -102.647
21.698 | -102.6611
21.7016 | -102.6764
21.7053 | -102.6918
21.709 | -102.7058
21.7127 | -102.7203
21.7163 | -102.7356
21.72 | -102.7538
21.7237 | -102.7716
21.7274 | -102.7873
21.731 | -102.8045
21.7347 | -102.8208
21.7384 | -102.8369
21.742 | -102.8534
21.7457 | -102.8704
21.7494 | -102.8892
21.7531 | -102.9069
21.7567 | -102.9228
21.7604 | -102.9374
21.7641 | -102.9491
21.7678 | -102.9618
21.7714 | -102.9777
21.7751 | -102.9903
21.7788 | -103.003
21.7825 | -103.015
21.7861 | -103.0265
21.7898 | -103.0394
21.7935 | -103.0521
21.7971 | -103.065
21.8008 | -103.0768
21.8045 | -103.0866
21.8082 | -103.0975
21.8118 | -103.1081
21.8155 | -103.1168
21.8192 | -103.1265
21.8229 | -103.1349
21.8265 | -103.1427
21.8302 | -103.1495
21.8339 | -103.1587
21.8376 | -103.1668
21.8412 | -103.175
21.8449 | -103.1813
21.8486 | -103.1885
21.8522 | -103.1935
21.8559 | -103.199
21.8596 | -103.2021
21.8633 | -103.2057
21.8669 | -103.2106
21.8706 | -103.2162
21.8743 | -103.2207
21.878 | -103.2263
21.8816 | -103.2333
21.8853 | -103.2361
21.889 | -103.2371
21.8926 | -103.2399
21.8963 | -103.2437
21.9 | -103.2477
21.9037 | -103.2507
21.9073 | -103.2526
21.911 | -103.2545
21.9147 | -103.257
21.9184 | -103.2599
21.922 | -103.2619
21.9257 | -103.2662
21.9294 | -103.2693
21.9331 | -103.273
21.9367 | -103.2777
21.9404 | -103.2831
21.9441 | -103.2881
21.9477 | -103.2935
21.9514 | -103.2994
21.9551 | -103.3044
21.9588 | -103.308
21.9624 | -103.3146
21.9661 | -103.3237
21.9698 | -103.3333
21.9735 | -103.3421
21.9771 | -103.3498
21.9808 | -103.3564
21.9845 | -103.3638
21.9882 | -103.3701
21.9918 | -103.3772
21.9955 | -103.3854
21.9992 | -103.3937
22.0028 | -103.4033
22.0065 | -103.4131
22.0102 | -103.4223
22.0139 | -103.431
22.0175 | -103.4402
22.0212 | -103.4477
22.0249 | -103.4542
22.0286 | -103.4591
22.0322 | -103.4655
22.0359 | -103.4727
22.0396 | -103.4788
22.0433 | -103.4827
22.0469 | -103.4858
22.0506 | -103.488
22.0543 | -103.4899
22.0579 | -103.4925
22.0616 | -103.4935
22.0653 | -103.4935
22.069 | -103.494
22.0726 | -103.4942
22.0763 | -103.4939
22.08 | -103.4912
22.0837 | -103.4862
22.0873 | -103.4837
22.091 | -103.4806
22.0947 | -103.4761
22.0984 | -103.4711
22.102 | -103.4652
22.1057 | -103.4589
22.1094 | -103.4543
22.113 | -103.4495
22.1167 | -103.4446
22.1204 | -103.4394
22.1241 | -103.4344
22.1277 | -103.4308
22.1314 | -103.4268
22.1351 | -103.4217
22.1388 | -103.4161
22.1424 | -103.4113
22.1461 | -103.4068
22.1498 | -103.4048
22.1535 | -103.4029
22.1571 | -103.4015
22.1608 | -103.4018
22.1645 | -103.4019
22.1681 | -103.4004
22.1718 | -103.4035
22.1755 | -103.4058
22.1792 | -103.408
22.1828 | -103.4105
22.1865 | -103.4152
22.1902 | -103.4225
22.1939 | -103.4296
22.1975 | -103.4367
22.2012 | -103.4461
22.2049 | -103.4554
22.2085 | -103.4638
22.2122 | -103.4724
22.2159 | -103.4824
22.2196 | -103.4927
22.2232 | -103.5059
22.2269 | -103.5186
22.2306 | -103.5314
22.2343 | -103.5454
22.2379 | -103.559
22.2416 | -103.5727
22.2453 | -103.5882
22.249 | -103.6011
22.2526 | -103.6148
22.2563 | -103.6284
22.26 | -103.6415
22.2636 | -103.6543
22.2673 | -103.6666
22.271 | -103.6801
22.2747 | -103.6942
22.2783 | -103.7061
22.282 | -103.7173
22.2857 | -103.7288
22.2894 | -103.7403
22.293 | -103.7522
22.2967 | -103.7631
22.3004 | -103.7721
22.3041 | -103.7821
22.3077 | -103.79
22.3114 | -103.7956
22.3151 | -103.801
22.3187 | -103.8062
22.3224 | -103.8124
22.3261 | -103.8175
22.3298 | -103.8219
22.3334 | -103.8259
22.3371 | -103.8312
22.3408 | -103.8351
22.3445 | -103.8387
22.3481 | -103.8389
22.3518 | -103.8371
22.3555 | -103.8353
22.3592 | -103.8339
22.3628 | -103.8332
22.3665 | -103.8322
22.3702 | -103.8275
22.3738 | -103.817
22.3775 | -103.803
22.3812 | -103.79
22.3849 | -103.778
22.3885 | -103.7637
22.3922 | -103.7502
22.3959 | -103.7348
22.3996 | -103.7126
22.4032 | -103.6919
22.4069 | -103.6704
22.4106 | -103.6458
22.4143 | -103.6132
22.4179 | -103.5801
22.4216 | -103.5504
22.4253 | -103.5212
22.4289 | -103.4908
22.4326 | -103.4593
22.4363 | -103.4253
22.44 | -103.3905
22.4436 | -103.3545
22.4473 | -103.3188
22.451 | -103.2836
22.4547 | -103.2472
22.4583 | -103.2097
22.462 | -103.1714
22.4657 | -103.1321
22.4694 | -103.0924
22.473 | -103.0514
22.4767 | -103.0103
22.4804 | -102.9697
22.484 | -102.9292
22.4877 | -102.8877
22.4914 | -102.846
22.4951 | -102.8057
22.4987 | -102.767
22.5024 | -102.7294
22.5061 | -102.6915
22.5098 | -102.6555
22.5134 | -102.6227
22.5171 | -102.592
22.5208 | -102.5632
22.5245 | -102.5374
22.5281 | -102.5135
22.5318 | -102.4917
22.5355 | -102.4732
22.5391 | -102.456
22.5428 | -102.4415
22.5465 | -102.4289
22.5502 | -102.4207
22.5538 | -102.4149
22.5575 | -102.4131
22.5612 | -102.4125
22.5649 | -102.4143
22.5685 | -102.4188
22.5722 | -102.4274
22.5759 | -102.4399
22.5795 | -102.4552
22.5832 | -102.4729
22.5869 | -102.4924
22.5906 | -102.5137
22.5942 | -102.5378
22.5979 | -102.565
22.6016 | -102.5927
22.6053 | -102.6206
22.6089 | -102.6509
22.6126 | -102.6834
22.6163 | -102.7177
22.62 | -102.7515
22.6236 | -102.7855
22.6273 | -102.822
22.631 | -102.858
22.6346 | -102.8925
22.6383 | -102.9279
22.642 | -102.9629
22.6457 | -102.9976
22.6493 | -103.0318
22.653 | -103.0637
22.6567 | -103.0925
22.6604 | -103.1239
22.664 | -103.1541
22.6677 | -103.1845
22.6714 | -103.2138
22.6751 | -103.2416
22.6787 | -103.2679
22.6824 | -103.2907
22.6861 | -103.314
22.6897 | -103.3388
22.6934 | -103.3643
22.6971 | -103.3896
22.7008 | -103.4112
22.7044 | -103.428
22.7081 | -103.4455
22.7118 | -103.4611
22.7155 | -103.4771
22.7191 | -103.4967
22.7228 | -103.5094
22.7265 | -103.5196
22.7302 | -103.5306
22.7338 | -103.5461
22.7375 | -103.5597
22.7412 | -103.5675
22.7448 | -103.5703
22.7485 | -103.5724
22.7522 | -103.5732
22.7559 | -103.5751
22.7595 | -103.5774
22.7632 | -103.5775
22.7669 | -103.5773
22.7706 | -103.5766
22.7742 | -103.5753
22.7779 | -103.5748
22.7816 | -103.5724
22.7853 | -103.5672
22.7889 | -103.5629
22.7926 | -103.5584
22.7963 | -103.555
22.7999 | -103.552
22.8036 | -103.549
22.8073 | -103.5443
22.811 | -103.5385
22.8146 | -103.5339
22.8183 | -103.5286
22.822 | -103.5219
22.8257 | -103.5155
22.8293 | -103.5097
22.833 | -103.5029
22.8367 | -103.4951
22.8404 | -103.4865
22.844 | -103.4784
22.8477 | -103.4697
22.8514 | -103.4602
22.855 | -103.4523
22.8587 | -103.444
22.8624 | -103.4361
22.8661 | -103.4296
22.8697 | -103.4225
22.8734 | -103.4142
22.8771 | -103.4065
22.8808 | -103.3979
22.8844 | -103.3904
22.8881 | -103.3836
22.8918 | -103.3746
22.8955 | -103.3656
22.8991 | -103.3566
22.9028 | -103.3476
22.9065 | -103.341
22.9101 | -103.3358
22.9138 | -103.3289
22.9175 | -103.3205
22.9212 | -103.3129
22.9248 | -103.3078
22.9285 | -103.3034
22.9322 | -103.2984
22.9359 | -103.2923
22.9395 | -103.2866
22.9432 | -103.2829
22.9469 | -103.279
22.9505 | -103.274
22.9542 | -103.2707
22.9579 | -103.2685
22.9616 | -103.2651
22.9652 | -103.2601
22.9689 | -103.2558
22.9726 | -103.2525
22.9763 | -103.251
22.9799 | -103.2509
22.9836 | -103.2476
22.9873 | -103.244
22.991 | -103.2393
22.9946 | -103.2349
22.9983 | -103.2313
23.002 | -103.2283
23.0056 | -103.2256
23.0093 | -103.2223
23.013 | -103.2182
23.0167 | -103.2137
23.0203 | -103.2099
23.024 | -103.2056
23.0277 | -103.2036
23.0314 | -103.2008
23.035 | -103.1981
23.0387 | -103.1944
23.0424 | -103.1915
23.0461 | -103.1898
23.0497 | -103.1874
23.0534 | -103.1854
23.0571 | -103.1831
23.0607 | -103.1791
23.0644 | -103.1738
23.0681 | -103.1705
23.0718 | -103.1679
23.0754 | -103.1656
23.0791 | -103.1624
23.0828 | -103.1577
23.0865 | -103.1554
23.0901 | -103.1526
23.0938 | -103.1502
23.0975 | -103.1491
23.1012 | -103.1473
23.1048 | -103.1465
23.1085 | -103.1446
23.1122 | -103.1414
23.1158 | -103.1407
23.1195 | -103.141
23.1232 | -103.1402
23.1269 | -103.1373
23.1305 | -103.1352
23.1342 | -103.1311
23.1379 | -103.1283
23.1416 | -103.127
23.1452 | -103.1238
23.1489 | -103.1179
23.1526 | -103.1124
23.1563 | -103.1077
23.1599 | -103.1027
23.1636 | -103.0985
23.1673 | -103.094
23.1709 | -103.0876
23.1746 | -103.0813
23.1783 | -103.0741
23.182 | -103.0651
23.1856 | -103.0572
23.1893 | -103.0491
23.193 | -103.0409
23.1967 | -103.0335
23.2003 | -103.0257
23.204 | -103.0175
23.2077 | -103.0073
23.2113 | -102.9958
23.215 | -102.9857
23.2187 | -102.9781
23.2224 | -102.9696
23.226 | -102.9597
23.2297 | -102.9491
23.2334 | -102.9366
23.2371 | -102.9252
23.2407 | -102.9161
23.2444 | -102.9086
23.2481 | -102.899
23.2518 | -102.8869
23.2554 | -102.8758
23.2591 | -102.8651
23.2628 | -102.8562
23.2664 | -102.846
23.2701 | -102.8352
23.2738 | -102.8241
23.2775 | -102.8143
23.2811 | -102.8052
23.2848 | -102.7951
23.2885 | -102.7853
23.2922 | -102.7776
23.2958 | -102.7688
23.2995 | -102.7599
23.3032 | -102.7507
23.3069 | -102.742
23.3105 | -102.7326
23.3142 | -102.7301
23.3179 | -102.7238
23.3215 | -102.7183
23.3252 | -102.7134
23.3289 | -102.7086
23.3326 | -102.7032
23.3362 | -102.6977
23.3399 | -102.6922
23.3436 | -102.6887
23.3473 | -102.6853
23.3509 | -102.6798
23.3546 | -102.6764
23.3583 | -102.6728
23.362 | -102.6696
23.3656 | -102.6665
23.3693 | -102.6649
23.373 | -102.6635
23.3766 | -102.6622
23.3803 | -102.6612
23.384 | -102.6603
23.3877 | -102.6613
23.3913 | -102.6651
23.395 | -102.6685
23.3987 | -102.6709
23.4024 | -102.6727
23.406 | -102.6756
23.4097 | -102.6793
23.4134 | -102.6843
23.4171 | -102.6894
23.4207 | -102.6948
23.4244 | -102.6997
23.4281 | -102.7031
23.4317 | -102.7077
23.4354 | -102.7136
23.4391 | -102.7196
23.4428 | -102.7255
23.4464 | -102.7299
23.4501 | -102.7337
23.4538 | -102.7381
23.4575 | -102.7436
23.4611 | -102.7494
23.4648 | -102.757
23.4685 | -102.763
23.4722 | -102.7679
23.4758 | -102.7748
23.4795 | -102.7819
23.4832 | -102.7883
23.4868 | -102.7941
23.4905 | -102.8001
23.4942 | -102.8047
23.4979 | -102.8094
23.5015 | -102.8155
23.5052 | -102.8213
23.5089 | -102.827
23.5126 | -102.834
23.5162 | -102.84
23.5199 | -102.8452
23.5236 | -102.8509
23.5273 | -102.8567
23.5309 | -102.8614
23.5346 | -102.8659
23.5383 | -102.8732
23.5419 | -102.8806
23.5456 | -102.8864
23.5493 | -102.8907
23.553 | -102.8938
23.5566 | -102.8963
23.5603 | -102.9013
23.564 | -102.9048
23.5677 | -102.9072
23.5713 | -102.9096
23.575 | -102.9113
23.5787 | -102.9132
23.5824 | -102.9172
23.586 | -102.921
23.5897 | -102.9245
23.5934 | -102.9273
23.597 | -102.9308
23.6007 | -102.9335
23.6044 | -102.9356
23.6081 | -102.937
23.6117 | -102.9382
23.6154 | -102.9418
23.6191 | -102.9445
23.6228 | -102.9475
23.6264 | -102.9513
23.6301 | -102.9542
23.6338 | -102.9572
23.6374 | -102.9598
23.6411 | -102.9631
23.6448 | -102.9647
23.6485 | -102.9639
23.6521 | -102.9632
23.6558 | -102.9626
23.6595 | -102.9636
23.6632 | -102.965
23.6668 | -102.9633
23.6705 | -102.9623
23.6742 | -102.9622
23.6779 | -102.9619
23.6815 | -102.9605
23.6852 | -102.9594
23.6889 | -102.9563
23.6925 | -102.9545
23.6962 | -102.9519
23.6999 | -102.9462
23.7036 | -102.9387
23.7072 | -102.9313
23.7109 | -102.9262
23.7146 | -102.921
23.7183 | -102.9166
23.7219 | -102.9109
23.7256 | -102.9027
23.7293 | -102.8955
23.733 | -102.889
23.7366 | -102.8815
23.7403 | -102.8721
23.744 | -102.8627
23.7476 | -102.8529
23.7513 | -102.8428
23.755 | -102.8332
23.7587 | -102.8233
23.7623 | -102.8145
23.766 | -102.8043
23.7697 | -102.7942
23.7734 | -102.784
23.777 | -102.7749
23.7807 | -102.768
23.7844 | -102.7613
23.7881 | -102.7552
23.7917 | -102.7496
23.7954 | -102.743
23.7991 | -102.7371
23.8027 | -102.7337
23.8064 | -102.7307
23.8101 | -102.7274
23.8138 | -102.7258
23.8174 | -102.7241
23.8211 | -102.7226
23.8248 | -102.7245
23.8285 | -102.7262
23.8321 | -102.7287
23.8358 | -102.7303
23.8395 | -102.7337
23.8432 | -102.7384
23.8468 | -102.7473
23.8505 | -102.7524
23.8542 | -102.7589
23.8578 | -102.7649
23.8615 | -102.7708
23.8652 | -102.7734
23.8689 | -102.7763
23.8725 | -102.7806
23.8762 | -102.7849
23.8799 | -102.7914
23.8836 | -102.7986
23.8872 | -102.8051
23.8909 | -102.8111
23.8946 | -102.8182
23.8983 | -102.8239
23.9019 | -102.8297
23.9056 | -102.8352
23.9093 | -102.8387
23.9129 | -102.842
23.9166 | -102.8443
23.9203 | -102.8443
23.924 | -102.8438
23.9276 | -102.8423
23.9313 | -102.8393
23.935 | -102.8356
23.9387 | -102.8342
23.9423 | -102.8302
23.946 | -102.8232
23.9497 | -102.8171
23.9534 | -102.8097
23.957 | -102.801
23.9607 | -102.7911
23.9644 | -102.7827
23.968 | -102.7728
23.9717 | -102.7614
23.9754 | -102.7501
23.9791 | -102.7402
23.9827 | -102.7283
23.9864 | -102.7168
23.9901 | -102.7074
23.9938 | -102.6979
23.9974 | -102.6861
24.0011 | -102.6736
24.0048 | -102.6637
24.0084 | -102.6531
24.0121 | -102.6404
24.0158 | -102.6278
24.0195 | -102.6162
24.0231 | -102.6078
24.0268 | -102.5984
24.0305 | -102.5875
24.0342 | -102.579
24.0378 | -102.5723
24.0415 | -102.5676
24.0452 | -102.5608
24.0489 | -102.553
24.0525 | -102.5466
24.0562 | -102.5393
24.0599 | -102.5326
24.0635 | -102.5278
24.0672 | -102.525
24.0709 | -102.5226
24.0746 | -102.5214
24.0782 | -102.5215
24.0819 | -102.5223
24.0856 | -102.5235
24.0893 | -102.5253
24.0929 | -102.5266
24.0966 | -102.5294
24.1003 | -102.5324
24.104 | -102.5333
24.1076 | -102.5349
24.1113 | -102.5395
24.115 | -102.5423
24.1186 | -102.5438
24.1223 | -102.5451
24.126 | -102.5465
24.1297 | -102.5458
24.1333 | -102.5451
24.137 | -102.5452
24.1407 | -102.5437
24.1444 | -102.5431
24.148 | -102.5423
24.1517 | -102.5416
24.1554 | -102.5393
24.1591 | -102.5365
24.1627 | -102.5316
24.1664 | -102.5276
24.1701 | -102.5237
24.1737 | -102.5198
24.1774 | -102.5124
24.1811 | -102.5051
24.1848 | -102.4989
24.1884 | -102.4913
24.1921 | -102.4823
24.1958 | -102.4706
24.1995 | -102.4605
24.2031 | -102.4487
24.2068 | -102.4399
24.2105 | -102.4325
24.2142 | -102.4221
24.2178 | -102.4091
24.2215 | -102.3941
24.2252 | -102.3794
24.2288 | -102.3652
24.2325 | -102.352
24.2362 | -102.3378
24.2399 | -102.3228
24.2435 | -102.3096
24.2472 | -102.2947
24.2509 | -102.2804
24.2546 | -102.268
24.2582 | -102.2525
24.2619 | -102.2216
24.2656 | -102.1292
24.2692 | -101.9962
24.2729 | -101.8674
24.2766 | -101.727
24.2803 | -101.6084
24.2839 | -101.5044
24.2876 | -101.4132
24.2913 | -101.322
24.295 | -101.2283
24.2986 | -101.1317
24.3023 | -101.0304
24.306 | -100.924
24.3097 | -100.8144
24.3133 | -100.7012
24.317 | -100.5847
24.3207 | -100.4645
24.3243 | -100.3419
24.328 | -100.217
24.3317 | -100.0897
24.3354 | -99.9615
24.339 | -99.8319
24.3427 | -99.702
24.3464 | -99.5716
24.3501 | -99.4406
24.3537 | -99.3103
24.3574 | -99.1813
24.3611 | -99.0549
24.3648 | -98.9299
24.3684 | -98.8079
24.3721 | -98.6896
24.3758 | -98.5845
24.3794 | -98.4775
24.3831 | -98.3753
24.3868 | -98.2767
24.3905 | -98.1841
24.3941 | -98.0966
24.3978 | -98.0149
24.4015 | -97.9385
24.4052 | -97.8677
24.4088 | -97.8029
24.4125 | -97.7448
24.4162 | -97.6936
24.4199 | -97.6487
24.4235 | -97.6103
24.4272 | -97.5789
24.4309 | -97.5548
24.4345 | -97.5372
24.4382 | -97.5262
24.4419 | -97.5224
24.4456 | -97.5261
24.4492 | -97.5377
24.4529 | -97.5561
24.4566 | -97.5805
24.4603 | -97.6125
24.4639 | -97.6512
24.4676 | -97.697
24.4713 | -97.7493
24.475 | -97.8082
24.4786 | -97.8743
24.4823 | -97.9465
24.486 | -98.0244
24.4896 | -98.1071
24.4933 | -98.1958
24.497 | -98.2896
24.5007 | -98.3888
24.5043 | -98.4921
24.508 | -98.6006
24.5117 | -98.7131
24.5154 | -98.8294
24.519 | -98.9486
24.5227 | -99.0717
24.5264 | -99.1969
24.5301 | -99.3225
24.5337 | -99.4496
24.5374 | -99.578
24.5411 | -99.7052
24.5447 | -99.8298
24.5484 | -99.9438
24.5521 | -100.0447
24.5558 | -100.1281
24.5594 | -100.1967
24.5631 | -100.2691
24.5668 | -100.333
24.5705 | -100.3907
24.5741 | -100.4285
24.5778 | -100.4599
24.5815 | -100.489
24.5852 | -100.5118
24.5888 | -100.5269
24.5925 | -100.5347
24.5962 | -100.533
24.5998 | -100.5196
24.6035 | -100.4957
24.6072 | -100.4693
24.6109 | -100.4661
24.6145 | -100.4702
24.6182 | -100.4496
24.6219 | -100.4169
24.6256 | -100.3488
24.6292 | -100.259
24.6329 | -100.1514
24.6366 | -100.0359
24.6403 | -99.9176
24.6439 | -99.7966
24.6476 | -99.6764
24.6513 | -99.5576
24.6549 | -99.4404
24.6586 | -99.3232
24.6623 | -99.2071
24.666 | -99.0938
24.6696 | -98.9841
24.6733 | -98.877
24.677 | -98.7722
24.6807 | -98.6697
24.6843 | -98.5727
24.688 | -98.4806
24.6917 | -98.3932
24.6953 | -98.3113
24.699 | -98.2351
24.7027 | -98.1637
24.7064 | -98.0978
24.71 | -98.0378
24.7137 | -97.9844
24.7174 | -97.937
24.7211 | -97.8954
24.7247 | -97.8607
24.7284 | -97.8326
24.7321 | -97.8108
24.7358 | -97.796
24.7394 | -97.7874
24.7431 | -97.7853
24.7468 | -97.7896
24.7504 | -97.8005
24.7541 | -97.8176
24.7578 | -97.8413
24.7615 | -97.8712
24.7651 | -97.9063
24.7688 | -97.9483
24.7725 | -97.997
24.7762 | -98.0508
24.7798 | -98.1098
24.7835 | -98.1747
24.7872 | -98.2455
24.7909 | -98.3211
24.7945 | -98.4011
24.7982 | -98.4854
24.8019 | -98.5739
24.8055 | -98.6663
24.8092 | -98.7625
24.8129 | -98.8619
24.8166 | -98.9633
24.8202 | -99.0674
24.8239 | -99.1748
24.8276 | -99.2822
24.8313 | -99.3913
24.8349 | -99.5023
24.8386 | -99.6152
24.8423 | -99.7278
24.846 | -99.8396
24.8496 | -99.9498
24.8533 | -100.0609
24.857 | -100.1694
24.8606 | -100.2759
24.8643 | -100.3805
24.868 | -100.4844
24.8717 | -100.5845
24.8753 | -100.6811
24.879 | -100.7773
24.8827 | -100.8705
24.8864 | -100.9602
24.89 | -101.05
24.8937 | -101.138
24.8974 | -101.2321
24.9011 | -101.3277
24.9047 | -101.3504
24.9084 | -101.4496
24.9121 | -101.5506
24.9157 | -101.6295
24.9194 | -101.7005
24.9231 | -101.7588
24.9268 | -101.8233
24.9304 | -101.8729
24.9341 | -101.9049
24.9378 | -101.9289
24.9415 | -101.9455
24.9451 | -101.9562
24.9488 | -101.9632
24.9525 | -101.9695
24.9562 | -101.9774
24.9598 | -101.9854
24.9635 | -101.9922
24.9672 | -101.9937
24.9708 | -101.9932
24.9745 | -101.9908
24.9782 | -101.9884
24.9819 | -101.9848
24.9855 | -101.9804
24.9892 | -101.977
24.9929 | -101.9723
24.9966 | -101.9691
25.0002 | -101.9642
25.0039 | -101.9577
25.0076 | -101.9502
25.0112 | -101.9451
25.0149 | -101.9414
25.0186 | -101.9356
25.0223 | -101.9286
25.0259 | -101.9236
25.0296 | -101.9213
25.0333 | -101.9195
25.037 | -101.9165
25.0406 | -101.9144
25.0443 | -101.9138
25.048 | -101.9141
25.0517 | -101.9142
25.0553 | -101.915
25.059 | -101.9156
25.0627 | -101.9178
25.0663 | -101.9204
25.07 | -101.9248
25.0737 | -101.9302
25.0774 | -101.9363
25.081 | -101.9441
25.0847 | -101.953
25.0884 | -101.9636
25.0921 | -101.9756
25.0957 | -101.9877
25.0994 | -102.0013
25.1031 | -102.017
25.1068 | -102.0337
25.1104 | -102.0509
25.1141 | -102.0673
25.1178 | -102.0867
25.1214 | -102.1064
25.1251 | -102.1247
25.1288 | -102.1447
25.1325 | -102.1652
25.1361 | -102.1856
25.1398 | -102.2068
25.1435 | -102.2291
25.1472 | -102.2511
25.1508 | -102.2725
25.1545 | -102.2938
25.1582 | -102.317
25.1619 | -102.3391
25.1655 | -102.36
25.1692 | -102.382
25.1729 | -102.404
25.1765 | -102.425
25.1802 | -102.4458
25.1839 | -102.4643
25.1876 | -102.4818
25.1912 | -102.4978
25.1949 | -102.5126
25.1986 | -102.528
25.2023 | -102.5449
25.2059 | -102.5591
25.2096 | -102.5729
25.2133 | -102.5873
25.217 | -102.5993
25.2206 | -102.6104
25.2243 | -102.6226
25.228 | -102.6353
25.2316 | -102.6457
25.2353 | -102.6541
25.239 | -102.662
25.2427 | -102.6685
25.2463 | -102.6747
25.25 | -102.6833
25.2537 | -102.6953
25.2574 | -102.7036
25.261 | -102.7081
25.2647 | -102.7169
25.2684 | -102.7234
25.2721 | -102.7262
25.2757 | -102.7292
25.2794 | -102.7333
25.2831 | -102.7369
25.2867 | -102.7431
25.2904 | -102.7493
25.2941 | -102.7528
25.2978 | -102.7557
25.3014 | -102.7592
25.3051 | -102.7615
25.3088 | -102.7597
25.3125 | -102.7605
25.3161 | -102.7617
25.3198 | -102.7613
25.3235 | -102.7626
25.3271 | -102.765
25.3308 | -102.7657
25.3345 | -102.7658
25.3382 | -102.7668
25.3418 | -102.7679
25.3455 | -102.7673
25.3492 | -102.7671
25.3529 | -102.7684
25.3565 | -102.7703
25.3602 | -102.7713
25.3639 | -102.7719
25.3676 | -102.7733
25.3712 | -102.7719
25.3749 | -102.7722
25.3786 | -102.7742
25.3822 | -102.7749
25.3859 | -102.7736
25.3896 | -102.7734
25.3933 | -102.7698
25.3969 | -102.7629
25.4006 | -102.7572
25.4043 | -102.7513
25.408 | -102.7472
25.4116 | -102.743
25.4153 | -102.7392
25.419 | -102.735
25.4227 | -102.7292
25.4263 | -102.7221
25.43 | -102.7171
25.4337 | -102.7111
25.4373 | -102.7067
25.441 | -102.7025
25.4447 | -102.6976
25.4484 | -102.6918
25.452 | -102.6869
25.4557 | -102.6839
25.4594 | -102.68
25.4631 | -102.676
25.4667 | -102.6708
25.4704 | -102.664
25.4741 | -102.6603
25.4778 | -102.6578
25.4814 | -102.6564
25.4851 | -102.6548
25.4888 | -102.6542
25.4924 | -102.653
25.4961 | -102.6496
25.4998 | -102.6467
25.5035 | -102.6439
25.5071 | -102.6433
25.5108 | -102.6434
25.5145 | -102.6429
25.5182 | -102.6416
25.5218 | -102.6414
25.5255 | -102.6432
25.5292 | -102.6453
25.5329 | -102.6469
25.5365 | -102.6491
25.5402 | -102.6518
25.5439 | -102.6553
25.5475 | -102.6595
25.5512 | -102.6628
25.5549 | -102.6672
25.5586 | -102.6743
25.5622 | -102.6796
25.5659 | -102.6854
25.5696 | -102.6896
25.5733 | -102.6917
25.5769 | -102.6952
25.5806 | -102.6995
25.5843 | -102.7042
25.588 | -102.7084
25.5916 | -102.7122
25.5953 | -102.7164
25.599 | -102.7209
25.6026 | -102.725
25.6063 | -102.7313
25.61 | -102.7369
25.6137 | -102.7432
25.6173 | -102.7497
25.621 | -102.7552
25.6247 | -102.7589
25.6284 | -102.7638
25.632 | -102.7681
25.6357 | -102.7725
25.6394 | -102.7763
25.6431 | -102.7798
25.6467 | -102.7819
25.6504 | -102.7832
25.6541 | -102.7856
25.6577 | -102.7888
25.6614 | -102.7906
25.6651 | -102.7938
25.6688 | -102.7985
25.6724 | -102.7988
25.6761 | -102.7994
25.6798 | -102.8018
25.6835 | -102.8013
25.6871 | -102.7996
25.6908 | -102.7991
25.6945 | -102.7997
25.6981 | -102.7985
25.7018 | -102.797
25.7055 | -102.7961
25.7092 | -102.7949
25.7128 | -102.7915
25.7165 | -102.7884
25.7202 | -102.785
25.7239 | -102.7829
25.7275 | -102.7821
25.7312 | -102.782
25.7349 | -102.7802
25.7386 | -102.7774
25.7422 | -102.7756
25.7459 | -102.7743
25.7496 | -102.7738
25.7532 | -102.7747
25.7569 | -102.7763
25.7606 | -102.7768
25.7643 | -102.7777
25.7679 | -102.7768
25.7716 | -102.7752
25.7753 | -102.7741
25.779 | -102.773
25.7826 | -102.7706
25.7863 | -102.7686
25.79 | -102.7661
25.7937 | -102.7646
25.7973 | -102.7626
25.801 | -102.7601
25.8047 | -102.7557
25.8083 | -102.7521
25.812 | -102.7485
25.8157 | -102.7454
25.8194 | -102.7414
25.823 | -102.737
25.8267 | -102.7323
25.8304 | -102.7292
25.8341 | -102.7256
25.8377 | -102.7205
25.8414 | -102.7123
25.8451 | -102.7038
25.8488 | -102.697
25.8524 | -102.6895
25.8561 | -102.6824
25.8598 | -102.6775
25.8634 | -102.6719
25.8671 | -102.6653
25.8708 | -102.6561
25.8745 | -102.6475
25.8781 | -102.6401
25.8818 | -102.6333
25.8855 | -102.6243
25.8892 | -102.6161
25.8928 | -102.6072
25.8965 | -102.5974
25.9002 | -102.588
25.9039 | -102.5811
25.9075 | -102.5739
25.9112 | -102.5656
25.9149 | -102.5565
25.9185 | -102.5487
25.9222 | -102.5406
25.9259 | -102.5321
25.9296 | -102.5234
25.9332 | -102.514
25.9369 | -102.5052
25.9406 | -102.4969
25.9443 | -102.4887
25.9479 | -102.4813
25.9516 | -102.4741
25.9553 | -102.4687
25.959 | -102.4626
25.9626 | -102.4559
25.9663 | -102.4515
25.97 | -102.4453
25.9736 | -102.4383
25.9773 | -102.4318
25.981 | -102.4252
25.9847 | -102.4207
25.9883 | -102.4143
25.992 | -102.408
25.9957 | -102.403
25.9994 | -102.3991
26.003 | -102.3952
26.0067 | -102.3913
26.0104 | -102.3868
26.0141 | -102.3829
26.0177 | -102.3783
26.0214 | -102.3711
26.0251 | -102.3651
26.0287 | -102.3614
26.0324 | -102.3578
26.0361 | -102.3541
26.0398 | -102.3525
26.0434 | -102.3496
26.0471 | -102.3443
26.0508 | -102.3397
26.0545 | -102.337
26.0581 | -102.3342
26.0618 | -102.3313
26.0655 | -102.3276
26.0691 | -102.3236
26.0728 | -102.3204
26.0765 | -102.3173
26.0802 | -102.3165
26.0838 | -102.3146
26.0875 | -102.3124
26.0912 | -102.3096
26.0949 | -102.3052
26.0985 | -102.301
26.1022 | -102.2993
26.1059 | -102.2986
26.1096 | -102.2957
26.1132 | -102.2938
26.1169 | -102.2917
26.1206 | -102.2899
26.1242 | -102.289
26.1279 | -102.2873
26.1316 | -102.2848
26.1353 | -102.2822
26.1389 | -102.2817
26.1426 | -102.2809
26.1463 | -102.2789
26.15 | -102.2778
26.1536 | -102.2781
26.1573 | -102.2778
26.161 | -102.2763
26.1647 | -102.2755
26.1683 | -102.2751
26.172 | -102.2751
26.1757 | -102.2756
26.1793 | -102.2774
26.183 | -102.2767
26.1867 | -102.2761
26.1904 | -102.2762
26.194 | -102.2767
26.1977 | -102.2783
26.2014 | -102.28
26.2051 | -102.281
26.2087 | -102.2824
26.2124 | -102.2843
26.2161 | -102.2852
26.2198 | -102.2877
26.2234 | -102.2888
26.2271 | -102.2874
26.2308 | -102.2862
26.2344 | -102.2845
26.2381 | -102.2827
26.2418 | -102.2837
26.2455 | -102.2831
26.2491 | -102.2821
26.2528 | -102.2805
26.2565 | -102.2773
26.2602 | -102.276
26.2638 | -102.2754
26.2675 | -102.2731
26.2712 | -102.2691
26.2749 | -102.2648
26.2785 | -102.2624
26.2822 | -102.2589
26.2859 | -102.2556
26.2895 | -102.2523
26.2932 | -102.2479
26.2969 | -102.2425
26.3006 | -102.2377
26.3042 | -102.2348
26.3079 | -102.2303
26.3116 | -102.2249
26.3153 | -102.2192
26.3189 | -102.2139
26.3226 | -102.2049
26.3263 | -102.1976
26.33 | -102.1896
26.3336 | -102.1816
26.3373 | -102.1743
26.341 | -102.1681
26.3446 | -102.1603
26.3483 | -102.1497
26.352 | -102.1395
26.3557 | -102.1299
26.3593 | -102.1227
26.363 | -102.1155
26.3667 | -102.106
26.3704 | -102.0959
26.374 | -102.0861
26.3777 | -102.0763
26.3814 | -102.0671
26.385 | -102.0562
26.3887 | -102.0458
26.3924 | -102.0383
26.3961 | -102.0285
26.3997 | -102.018
26.4034 | -102.0054
26.4071 | -101.9951
26.4108 | -101.9844
26.4144 | -101.9741
26.4181 | -101.965
26.4218 | -101.9574
26.4255 | -101.949
26.4291 | -101.9409
26.4328 | -101.9315
26.4365 | -101.9239
26.4401 | -101.918
26.4438 | -101.9115
26.4475 | -101.9054
26.4512 | -101.8988
26.4548 | -101.8943
26.4585 | -101.8907
26.4622 | -101.8873
26.4659 | -101.8816
26.4695 | -101.8776
26.4732 | -101.8729
26.4769 | -101.8695
26.4806 | -101.8685
26.4842 | -101.8673
26.4879 | -101.8671
26.4916 | -101.8672
26.4952 | -101.8731
26.4989 | -101.8728
26.5026 | -101.8748
26.5063 | -101.8765
26.5099 | -101.8778
26.5136 | -101.8788
26.5173 | -101.8819
26.521 | -101.8868
26.5246 | -101.8919
26.5283 | -101.8951
26.532 | -101.8984
26.5357 | -101.903
26.5393 | -101.9089
26.543 | -101.9135
26.5467 | -101.919
26.5503 | -101.925
26.554 | -101.9298
26.5577 | -101.9352
26.5614 | -101.9415
26.565 | -101.9469
26.5687 | -101.9524
26.5724 | -101.9582
26.5761 | -101.9634
26.5797 | -101.9694
26.5834 | -101.9771
26.5871 | -101.9828
26.5908 | -101.9886
26.5944 | -101.9931
26.5981 | -101.9982
26.6018 | -102.007
26.6054 | -102.0139
26.6091 | -102.0202
26.6128 | -102.0272
26.6165 | -102.0319
26.6201 | -102.0375
26.6238 | -102.0438
26.6275 | -102.0503
26.6312 | -102.0576
26.6348 | -102.0647
26.6385 | -102.0715
26.6422 | -102.0782
26.6459 | -102.0842
26.6495 | -102.0896
26.6532 | -102.0948
26.6569 | -102.1007
26.6605 | -102.1051
26.6642 | -102.1106
26.6679 | -102.1178
26.6716 | -102.1252
26.6752 | -102.1321
26.6789 | -102.1361
26.6826 | -102.1423
26.6863 | -102.1481
26.6899 | -102.1516
26.6936 | -102.1557
26.6973 | -102.1627
26.7009 | -102.1691
26.7046 | -102.1748
26.7083 | -102.1794
26.712 | -102.1824
26.7156 | -102.1843
26.7193 | -102.1876
26.723 | -102.1932
26.7267 | -102.1984
26.7303 | -102.2011
26.734 | -102.204
26.7377 | -102.2067
26.7414 | -102.2065
26.745 | -102.2069
26.7487 | -102.2098
26.7524 | -102.212
26.756 | -102.2134
26.7597 | -102.2152
26.7634 | -102.2173
26.7671 | -102.2153
26.7707 | -102.2127
26.7744 | -102.2105
26.7781 | -102.2116
26.7818 | -102.208
26.7854 | -102.2028
26.7891 | -102.1985
26.7928 | -102.1941
26.7965 | -102.1891
26.8001 | -102.1826
26.8038 | -102.1759
26.8075 | -102.1694
26.8111 | -102.1631
26.8148 | -102.1552
26.8185 | -102.148
26.8222 | -102.1407
26.8258 | -102.1325
26.8295 | -102.1222
26.8332 | -102.1139
26.8369 | -102.1045
26.8405 | -102.0966
26.8442 | -102.0876
26.8479 | -102.078
26.8516 | -102.0666
26.8552 | -102.0546
26.8589 | -102.044
26.8626 | -102.0327
26.8662 | -102.02
26.8699 | -102.007
26.8736 | -101.9954
26.8773 | -101.9832
26.8809 | -101.9701
26.8846 | -101.9594
26.8883 | -101.9499
26.892 | -101.9382
26.8956 | -101.924
26.8993 | -101.9107
26.903 | -101.8991
26.9067 | -101.8888
26.9103 | -101.8795
26.914 | -101.8698
26.9177 | -101.8607
26.9213 | -101.8533
26.925 | -101.8459
26.9287 | -101.8387
26.9324 | -101.8325
26.936 | -101.8257
26.9397 | -101.8189
26.9434 | -101.8136
26.9471 | -101.8078
26.9507 | -101.8013
26.9544 | -101.7956
26.9581 | -101.7892
26.9618 | -101.7843
26.9654 | -101.7831
26.9691 | -101.783
26.9728 | -101.7831
26.9764 | -101.7856
26.9801 | -101.7872
26.9838 | -101.79
26.9875 | -101.7932
26.9911 | -101.7956
26.9948 | -101.7989
26.9985 | -101.8021
27.0022 | -101.8056
27.0058 | -101.8098
27.0095 | -101.8143
27.0132 | -101.8192
27.0169 | -101.8251
27.0205 | -101.8328
27.0242 | -101.8415
27.0279 | -101.8484
27.0315 | -101.8556
27.0352 | -101.8628
27.0389 | -101.8718
27.0426 | -101.8799
27.0462 | -101.8851
27.0499 | -101.8893
27.0536 | -101.8936
27.0573 | -101.898
27.0609 | -101.9033
27.0646 | -101.9043
27.0683 | -101.9056
27.072 | -101.9086
27.0756 | -101.9119
27.0793 | -101.9157
27.083 | -101.9181
27.0866 | -101.922
27.0903 | -101.9262
27.094 | -101.9289
27.0977 | -101.9312
27.1013 | -101.9332
27.105 | -101.9345
27.1087 | -101.9352
27.1124 | -101.9349
27.116 | -101.9338
27.1197 | -101.9332
27.1234 | -101.9346
27.127 | -101.9338
27.1307 | -101.93
27.1344 | -101.9268
27.1381 | -101.9244
27.1417 | -101.9227
27.1454 | -101.9213
27.1491 | -101.9206
27.1528 | -101.9198
27.1564 | -101.9178
27.1601 | -101.9143
27.1638 | -101.9131
27.1675 | -101.9126
27.1711 | -101.9107
27.1748 | -101.9093
27.1785 | -101.908
27.1821 | -101.9065
27.1858 | -101.9035
27.1895 | -101.9012
27.1932 | -101.8967
27.1968 | -101.895
27.2005 | -101.8939
27.2042 | -101.8939
27.2079 | -101.8945
27.2115 | -101.8955
27.2152 | -101.8952
27.2189 | -101.897
27.2226 | -101.9007
27.2262 | -101.9046
27.2299 | -101.9079
27.2336 | -101.9121
27.2372 | -101.9181
27.2409 | -101.924
27.2446 | -101.9304
27.2483 | -101.9385
27.2519 | -101.9451
27.2556 | -101.9522
27.2593 | -101.9616
27.263 | -101.9715
27.2666 | -101.981
27.2703 | -101.9912
27.274 | -102.0034
27.2777 | -102.017
27.2813 | -102.03
27.285 | -102.043
27.2887 | -102.0557
27.2923 | -102.0672
27.296 | -102.0807
27.2997 | -102.0931
27.3034 | -102.1052
27.307 | -102.118
27.3107 | -102.1295
27.3144 | -102.1425
27.3181 | -102.1574
27.3217 | -102.172
27.3254 | -102.1874
27.3291 | -102.2039
27.3328 | -102.2187
27.3364 | -102.2345
27.3401 | -102.2504
27.3438 | -102.2627
27.3474 | -102.2729
27.3511 | -102.283
27.3548 | -102.2934
27.3585 | -102.303
27.3621 | -102.3114
27.3658 | -102.3198
27.3695 | -102.3297
27.3732 | -102.341
27.3768 | -102.3505
27.3805 | -102.3583
27.3842 | -102.366
27.3879 | -102.375
27.3915 | -102.3824
27.3952 | -102.389
27.3989 | -102.3944
27.4025 | -102.3982
27.4062 | -102.4035
27.4099 | -102.4099
27.4136 | -102.4162
27.4172 | -102.4216
27.4209 | -102.4251
27.4246 | -102.4277
27.4283 | -102.4313
27.4319 | -102.4351
27.4356 | -102.4394
27.4393 | -102.4428
27.4429 | -102.447
27.4466 | -102.4505
27.4503 | -102.4549
27.454 | -102.4582
27.4576 | -102.4605
27.4613 | -102.4615
27.465 | -102.4594
27.4687 | -102.4582
27.4723 | -102.4573
27.476 | -102.4556
27.4797 | -102.454
27.4834 | -102.4514
27.487 | -102.4492
27.4907 | -102.4467
27.4944 | -102.4437
27.498 | -102.4423
27.5017 | -102.4407
27.5054 | -102.4384
27.5091 | -102.4376
27.5127 | -102.4377
27.5164 | -102.4374
27.5201 | -102.4381
27.5238 | -102.4378
27.5274 | -102.4345
27.5311 | -102.4307
27.5348 | -102.4283
27.5385 | -102.4268
27.5421 | -102.4265
27.5458 | -102.4255
27.5495 | -102.424
27.5531 | -102.4219
27.5568 | -102.418
27.5605 | -102.4154
27.5642 | -102.4141
27.5678 | -102.4121
27.5715 | -102.4095
27.5752 | -102.4094
27.5789 | -102.4094
27.5825 | -102.411
27.5862 | -102.4086
27.5899 | -102.4044
27.5936 | -102.4002
27.5972 | -102.3969
27.6009 | -102.3942
27.6046 | -102.3901
27.6082 | -102.3857
27.6119 | -102.3801
27.6156 | -102.372
27.6193 | -102.363
27.6229 | -102.3555
27.6266 | -102.3473
27.6303 | -102.3389
27.634 | -102.3298
27.6376 | -102.3192
27.6413 | -102.3103
27.645 | -102.3011
27.6487 | -102.2918
27.6523 | -102.2811
27.656 | -102.2695
27.6597 | -102.2588
27.6633 | -102.2474
27.667 | -102.2349
27.6707 | -102.2231
27.6744 | -102.211
27.678 | -102.1981
27.6817 | -102.1856
27.6854 | -102.1726
27.6891 | -102.1594
27.6927 | -102.1461
27.6964 | -102.1317
27.7001 | -102.1169
27.7038 | -102.1019
27.7074 | -102.0866
27.7111 | -102.0703
27.7148 | -102.0511
27.7184 | -102.0328
27.7221 | -102.0175
27.7258 | -102.0016
27.7295 | -101.9837
27.7331 | -101.9677
27.7368 | -101.9524
27.7405 | -101.9374
27.7442 | -101.9223
27.7478 | -101.9082
27.7515 | -101.895
27.7552 | -101.8798
27.7588 | -101.8657
27.7625 | -101.8517
27.7662 | -101.8384
27.7699 | -101.8271
27.7735 | -101.8151
27.7772 | -101.8033
27.7809 | -101.7935
27.7846 | -101.783
27.7882 | -101.7733
27.7919 | -101.7649
27.7956 | -101.7579
27.7993 | -101.7489
27.8029 | -101.7413
27.8066 | -101.7358
27.8103 | -101.732
27.8139 | -101.7282
27.8176 | -101.7246
27.8213 | -101.7221
27.825 | -101.7214
27.8286 | -101.7222
27.8323 | -101.7227
27.836 | -101.7232
27.8397 | -101.7245
27.8433 | -101.7258
27.847 | -101.7282
27.8507 | -101.7309
27.8544 | -101.7363
27.858 | -101.7408
27.8617 | -101.7431
27.8654 | -101.7485
27.869 | -101.7561
27.8727 | -101.7647
27.8764 | -101.772
27.8801 | -101.7791
27.8837 | -101.7861
27.8874 | -101.7927
27.8911 | -101.8015
27.8948 | -101.8096
27.8984 | -101.8167
27.9021 | -101.824
27.9058 | -101.8316
27.9095 | -101.8381
27.9131 | -101.8435
27.9168 | -101.8505
27.9205 | -101.8584
27.9241 | -101.8637
27.9278 | -101.8603
27.9315 | -101.8538
27.9352 | -101.854
27.9388 | -101.8555
27.9425 | -101.8557
27.9462 | -101.8572
27.9499 | -101.8593
27.9535 | -101.8597
27.9572 | -101.8579
27.9609 | -101.856
27.9646 | -101.8532
27.9682 | -101.8494
27.9719 | -101.8443
27.9756 | -101.8378
27.9792 | -101.8326
27.9829 | -101.826
27.9866 | -101.8181
27.9903 | -101.8104
27.9939 | -101.8016
27.9976 | -101.7898
28.0013 | -101.7726