## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 260 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 249
Hood | 02-92B/P | 210
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1005.3 | hPa |  | PRC PRO 01
Humidity | 57.7 | %/Hr
Internal temperature of calibration bench | 23.76 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1010 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.1 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.01 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3105 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 04/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 10/07/2018
Technical Manager | Anthony HUE | Approval |  | 16/07/2017
CEO | Robert GLEMOT | Approval |  | 16/07/2017
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4209 | 57.4445
0.012209 | -35.9604 | 38.3603
0.018313 | -34.9109 | 28.4198
0.024417 | -34.4643 | 22.0538
0.030521 | -34.3598 | 18.2801
0.036626 | -34.2197 | 15.1875
0.04273 | -34.136 | 12.8762
0.048834 | -34.0846 | 11.3126
0.054938 | -34.0138 | 10.1159
0.061043 | -33.938 | 8.9732
0.067147 | -33.9274 | 8.0992
0.073251 | -33.9284 | 7.347
0.079355 | -33.9466 | 6.7934
0.08546 | -33.9647 | 6.3372
0.091564 | -33.9616 | 5.9594
0.097668 | -33.9284 | 5.7013
0.10377 | -33.8818 | 5.3625
0.10988 | -33.8515 | 5.0085
0.11598 | -33.8426 | 4.6574
0.12209 | -33.8592 | 4.3831
0.12819 | -33.8916 | 4.1495
0.13429 | -33.923 | 3.9056
0.1404 | -33.9406 | 3.6289
0.1465 | -33.9567 | 3.3664
0.15261 | -33.9644 | 3.1205
0.15871 | -33.9612 | 2.916
0.16482 | -33.9466 | 2.7501
0.17092 | -33.9219 | 2.685
0.17702 | -33.903 | 2.5853
0.18313 | -33.8858 | 2.5036
0.18923 | -33.8735 | 2.401
0.19534 | -33.8665 | 2.3034
0.20144 | -33.8651 | 2.1969
0.20754 | -33.8671 | 2.1285
0.21365 | -33.8689 | 2.0589
0.21975 | -33.8727 | 1.9973
0.22586 | -33.8801 | 1.9338
0.23196 | -33.8883 | 1.8613
0.23807 | -33.8967 | 1.7828
0.24417 | -33.899 | 1.689
0.25027 | -33.9025 | 1.6043
0.25638 | -33.9026 | 1.5126
0.26248 | -33.9011 | 1.4254
0.26859 | -33.8991 | 1.3425
0.27469 | -33.8987 | 1.2558
0.2808 | -33.9046 | 1.1907
0.2869 | -33.9091 | 1.1217
0.293 | -33.9146 | 1.0633
0.29911 | -33.9209 | 1.0235
0.30521 | -33.9269 | 0.9858
0.31132 | -33.931 | 0.97118
0.31742 | -33.9355 | 0.9484
0.32353 | -33.9391 | 0.93021
0.32963 | -33.9423 | 0.91608
0.33573 | -33.9444 | 0.90485
0.34184 | -33.9473 | 0.89051
0.34794 | -33.9507 | 0.85987
0.35405 | -33.9541 | 0.83329
0.36015 | -33.9577 | 0.8031
0.36626 | -33.9601 | 0.76922
0.37236 | -33.9625 | 0.72277
0.37846 | -33.9646 | 0.6683
0.38457 | -33.9632 | 0.60932
0.39067 | -33.9636 | 0.56047
0.39678 | -33.9617 | 0.5044
0.40288 | -33.9592 | 0.4506
0.40899 | -33.9551 | 0.39794
0.41509 | -33.9504 | 0.34525
0.42119 | -33.9454 | 0.29755
0.4273 | -33.9416 | 0.24771
0.4334 | -33.9381 | 0.19729
0.43951 | -33.9352 | 0.14806
0.44561 | -33.9338 | 0.10161
0.45172 | -33.9357 | 0.063414
0.45782 | -33.9356 | 0.021084
0.46392 | -33.9371 | -0.018226
0.47003 | -33.939 | -0.053773
0.47613 | -33.9413 | -0.086388
0.48224 | -33.9437 | -0.11742
0.48834 | -33.9461 | -0.13897
0.49445 | -33.9482 | -0.16422
0.50055 | -33.9507 | -0.18646
0.50665 | -33.9523 | -0.20508
0.51276 | -33.9534 | -0.22553
0.51886 | -33.9545 | -0.24316
0.52497 | -33.9533 | -0.25012
0.53107 | -33.9536 | -0.26058
0.53717 | -33.9538 | -0.27233
0.54328 | -33.9534 | -0.2786
0.54938 | -33.9532 | -0.28734
0.55549 | -33.9526 | -0.2968
0.56159 | -33.9525 | -0.30823
0.5677 | -33.9521 | -0.31906
0.5738 | -33.9515 | -0.33299
0.5799 | -33.9514 | -0.34512
0.58601 | -33.9519 | -0.3614
0.59211 | -33.9531 | -0.3837
0.59822 | -33.9538 | -0.40051
0.60432 | -33.9549 | -0.42265
0.61043 | -33.956 | -0.44377
0.61653 | -33.9573 | -0.46546
0.62263 | -33.9585 | -0.48843
0.62874 | -33.96 | -0.50894
0.63484 | -33.9612 | -0.53043
0.64095 | -33.9626 | -0.55205
0.64705 | -33.9639 | -0.57272
0.65316 | -33.9651 | -0.59266
0.65926 | -33.9662 | -0.61539
0.66536 | -33.9659 | -0.63443
0.67147 | -33.9663 | -0.65151
0.67757 | -33.967 | -0.67029
0.68368 | -33.9669 | -0.69098
0.68978 | -33.9671 | -0.70962
0.69589 | -33.9668 | -0.7287
0.70199 | -33.966 | -0.75159
0.70809 | -33.9655 | -0.77226
0.7142 | -33.9652 | -0.79227
0.7203 | -33.9644 | -0.81227
0.72641 | -33.9636 | -0.8337
0.73251 | -33.9634 | -0.85453
0.73862 | -33.9634 | -0.8783
0.74472 | -33.9636 | -0.89762
0.75082 | -33.9632 | -0.91845
0.75693 | -33.9632 | -0.94153
0.76303 | -33.9632 | -0.96339
0.76914 | -33.9635 | -0.98915
0.77524 | -33.9636 | -1.0136
0.78135 | -33.9639 | -1.0361
0.78745 | -33.9644 | -1.0596
0.79355 | -33.9647 | -1.0831
0.79966 | -33.9655 | -1.108
0.80576 | -33.9664 | -1.1369
0.81187 | -33.9666 | -1.1654
0.81797 | -33.967 | -1.195
0.82408 | -33.9674 | -1.2211
0.83018 | -33.9678 | -1.2512
0.83628 | -33.9684 | -1.2809
0.84239 | -33.9681 | -1.3087
0.84849 | -33.9679 | -1.3361
0.8546 | -33.9681 | -1.3635
0.8607 | -33.9679 | -1.3923
0.86681 | -33.9678 | -1.4205
0.87291 | -33.9671 | -1.4459
0.87901 | -33.9665 | -1.4693
0.88512 | -33.9663 | -1.4927
0.89122 | -33.966 | -1.5164
0.89733 | -33.9656 | -1.5359
0.90343 | -33.9652 | -1.5578
0.90953 | -33.9655 | -1.5738
0.91564 | -33.9648 | -1.5907
0.92174 | -33.9644 | -1.6056
0.92785 | -33.9638 | -1.62
0.93395 | -33.9638 | -1.6325
0.94006 | -33.9637 | -1.6468
0.94616 | -33.9637 | -1.6523
0.95226 | -33.9637 | -1.6644
0.95837 | -33.9639 | -1.6763
0.96447 | -33.9642 | -1.6871
0.97058 | -33.9642 | -1.697
0.97668 | -33.964 | -1.7048
0.98279 | -33.964 | -1.7074
0.98889 | -33.9643 | -1.7118
0.99499 | -33.9643 | -1.7203
1.0011 | -33.9646 | -1.7243
1.0072 | -33.9644 | -1.727
1.0133 | -33.9643 | -1.7297
1.0194 | -33.9642 | -1.7334
1.0255 | -33.9639 | -1.7396
1.0316 | -33.9632 | -1.7426
1.0377 | -33.9628 | -1.7473
1.0438 | -33.962 | -1.7502
1.0499 | -33.9605 | -1.7543
1.056 | -33.9593 | -1.7575
1.0621 | -33.9581 | -1.7618
1.0682 | -33.9576 | -1.7651
1.0743 | -33.9565 | -1.7712
1.0805 | -33.9558 | -1.7765
1.0866 | -33.9542 | -1.7855
1.0927 | -33.9528 | -1.7915
1.0988 | -33.9521 | -1.8008
1.1049 | -33.9506 | -1.807
1.111 | -33.9491 | -1.8167
1.1171 | -33.9477 | -1.8246
1.1232 | -33.9464 | -1.8384
1.1293 | -33.9453 | -1.8492
1.1354 | -33.9444 | -1.8609
1.1415 | -33.9432 | -1.8742
1.1476 | -33.9419 | -1.8875
1.1537 | -33.9406 | -1.9005
1.1598 | -33.9405 | -1.9192
1.1659 | -33.94 | -1.931
1.172 | -33.9394 | -1.944
1.1781 | -33.9387 | -1.9579
1.1842 | -33.9382 | -1.9735
1.1903 | -33.9391 | -1.9896
1.1964 | -33.939 | -2.0045
1.2025 | -33.939 | -2.0203
1.2086 | -33.9396 | -2.0363
1.2147 | -33.9399 | -2.0522
1.2209 | -33.9407 | -2.0661
1.227 | -33.9427 | -2.0821
1.2331 | -33.9437 | -2.0967
1.2392 | -33.9448 | -2.1131
1.2453 | -33.9461 | -2.1259
1.2514 | -33.9476 | -2.1389
1.2575 | -33.9492 | -2.1535
1.2636 | -33.9517 | -2.1654
1.2697 | -33.9536 | -2.1778
1.2758 | -33.9552 | -2.1924
1.2819 | -33.9571 | -2.2066
1.288 | -33.9588 | -2.221
1.2941 | -33.961 | -2.2319
1.3002 | -33.9631 | -2.2403
1.3063 | -33.9648 | -2.2508
1.3124 | -33.9667 | -2.2612
1.3185 | -33.9687 | -2.2724
1.3246 | -33.9707 | -2.2837
1.3307 | -33.9728 | -2.2934
1.3368 | -33.9746 | -2.3008
1.3429 | -33.9763 | -2.31
1.349 | -33.9778 | -2.3188
1.3551 | -33.9794 | -2.3296
1.3613 | -33.981 | -2.3375
1.3674 | -33.9822 | -2.3461
1.3735 | -33.9835 | -2.3545
1.3796 | -33.9847 | -2.3622
1.3857 | -33.9857 | -2.3693
1.3918 | -33.9866 | -2.3793
1.3979 | -33.9879 | -2.3878
1.404 | -33.9884 | -2.3979
1.4101 | -33.9892 | -2.4071
1.4162 | -33.99 | -2.4157
1.4223 | -33.9906 | -2.4261
1.4284 | -33.9913 | -2.4342
1.4345 | -33.9918 | -2.4447
1.4406 | -33.9921 | -2.456
1.4467 | -33.9926 | -2.4655
1.4528 | -33.9932 | -2.4749
1.4589 | -33.9938 | -2.4851
1.465 | -33.9939 | -2.4932
1.4711 | -33.9946 | -2.5023
1.4772 | -33.9949 | -2.515
1.4833 | -33.9954 | -2.5273
1.4894 | -33.9959 | -2.5391
1.4955 | -33.9965 | -2.551
1.5016 | -33.9971 | -2.5636
1.5078 | -33.9974 | -2.5789
1.5139 | -33.998 | -2.5923
1.52 | -33.9983 | -2.6062
1.5261 | -33.9986 | -2.6186
1.5322 | -33.9987 | -2.632
1.5383 | -33.999 | -2.645
1.5444 | -33.9996 | -2.66
1.5505 | -33.9999 | -2.6738
1.5566 | -34.0004 | -2.6876
1.5627 | -34.0009 | -2.7017
1.5688 | -34.0017 | -2.716
1.5749 | -34.0019 | -2.7291
1.581 | -34.0022 | -2.747
1.5871 | -34.0027 | -2.7637
1.5932 | -34.003 | -2.7777
1.5993 | -34.0034 | -2.7956
1.6054 | -34.0037 | -2.8118
1.6115 | -34.0041 | -2.8298
1.6176 | -34.0045 | -2.8475
1.6237 | -34.0047 | -2.8631
1.6298 | -34.005 | -2.8798
1.6359 | -34.0054 | -2.8955
1.642 | -34.0059 | -2.9112
1.6482 | -34.0062 | -2.9294
1.6543 | -34.0067 | -2.9473
1.6604 | -34.0072 | -2.9631
1.6665 | -34.0073 | -2.9792
1.6726 | -34.0076 | -2.9964
1.6787 | -34.0081 | -3.0146
1.6848 | -34.008 | -3.0316
1.6909 | -34.0083 | -3.049
1.697 | -34.0086 | -3.0661
1.7031 | -34.0086 | -3.0825
1.7092 | -34.0086 | -3.098
1.7153 | -34.0089 | -3.1149
1.7214 | -34.0085 | -3.1307
1.7275 | -34.0083 | -3.1471
1.7336 | -34.0081 | -3.1639
1.7397 | -34.0082 | -3.1793
1.7458 | -34.0083 | -3.1967
1.7519 | -34.0081 | -3.2132
1.758 | -34.008 | -3.2278
1.7641 | -34.0081 | -3.2425
1.7702 | -34.0079 | -3.2584
1.7763 | -34.008 | -3.272
1.7824 | -34.0079 | -3.2877
1.7885 | -34.0076 | -3.3011
1.7947 | -34.0078 | -3.3154
1.8008 | -34.0077 | -3.3278
1.8069 | -34.0076 | -3.3414
1.813 | -34.0076 | -3.3548
1.8191 | -34.0073 | -3.3687
1.8252 | -34.0072 | -3.3813
1.8313 | -34.0072 | -3.395
1.8374 | -34.0073 | -3.4077
1.8435 | -34.0074 | -3.419
1.8496 | -34.0071 | -3.4327
1.8557 | -34.007 | -3.4469
1.8618 | -34.007 | -3.4575
1.8679 | -34.007 | -3.4694
1.874 | -34.007 | -3.4814
1.8801 | -34.0069 | -3.4934
1.8862 | -34.0071 | -3.505
1.8923 | -34.007 | -3.516
1.8984 | -34.0067 | -3.5255
1.9045 | -34.0066 | -3.5356
1.9106 | -34.0063 | -3.5465
1.9167 | -34.0061 | -3.5567
1.9228 | -34.006 | -3.5664
1.9289 | -34.0057 | -3.5741
1.9351 | -34.0058 | -3.5821
1.9412 | -34.0058 | -3.5914
1.9473 | -34.0059 | -3.6023
1.9534 | -34.0058 | -3.6122
1.9595 | -34.0055 | -3.6225
1.9656 | -34.005 | -3.6321
1.9717 | -34.0048 | -3.6415
1.9778 | -34.0048 | -3.652
1.9839 | -34.0048 | -3.6623
1.99 | -34.0048 | -3.671
1.9961 | -34.0045 | -3.6805
2.0022 | -34.0045 | -3.6896
2.0083 | -34.0041 | -3.6985
2.0144 | -34.0041 | -3.7073
2.0205 | -34.0041 | -3.7181
2.0266 | -34.0042 | -3.7285
2.0327 | -34.0042 | -3.7386
2.0388 | -34.0044 | -3.7515
2.0449 | -34.0045 | -3.7611
2.051 | -34.0046 | -3.7701
2.0571 | -34.0044 | -3.7806
2.0632 | -34.0042 | -3.7903
2.0693 | -34.0042 | -3.8005
2.0754 | -34.0044 | -3.8117
2.0816 | -34.0046 | -3.8229
2.0877 | -34.0045 | -3.8335
2.0938 | -34.0047 | -3.8449
2.0999 | -34.0047 | -3.8559
2.106 | -34.0049 | -3.8678
2.1121 | -34.0049 | -3.8795
2.1182 | -34.0051 | -3.8891
2.1243 | -34.0053 | -3.8994
2.1304 | -34.0054 | -3.9092
2.1365 | -34.0057 | -3.9204
2.1426 | -34.0058 | -3.9334
2.1487 | -34.0062 | -3.9448
2.1548 | -34.0067 | -3.9575
2.1609 | -34.0073 | -3.9693
2.167 | -34.0077 | -3.9816
2.1731 | -34.0081 | -3.9933
2.1792 | -34.0088 | -4.0048
2.1853 | -34.0093 | -4.0174
2.1914 | -34.0097 | -4.0294
2.1975 | -34.0103 | -4.0413
2.2036 | -34.0108 | -4.0544
2.2097 | -34.0112 | -4.0663
2.2158 | -34.0121 | -4.077
2.222 | -34.0127 | -4.0875
2.2281 | -34.013 | -4.0984
2.2342 | -34.0135 | -4.1118
2.2403 | -34.014 | -4.122
2.2464 | -34.0148 | -4.1336
2.2525 | -34.0152 | -4.1448
2.2586 | -34.0162 | -4.1551
2.2647 | -34.017 | -4.1648
2.2708 | -34.0176 | -4.1768
2.2769 | -34.0185 | -4.1887
2.283 | -34.019 | -4.1989
2.2891 | -34.0195 | -4.2098
2.2952 | -34.02 | -4.22
2.3013 | -34.0207 | -4.2308
2.3074 | -34.0214 | -4.2409
2.3135 | -34.0219 | -4.2528
2.3196 | -34.0225 | -4.2621
2.3257 | -34.0229 | -4.2733
2.3318 | -34.0233 | -4.2829
2.3379 | -34.0238 | -4.2948
2.344 | -34.0241 | -4.3061
2.3501 | -34.0246 | -4.3174
2.3562 | -34.0249 | -4.3264
2.3623 | -34.0252 | -4.3355
2.3685 | -34.0254 | -4.3451
2.3746 | -34.0257 | -4.3563
2.3807 | -34.0262 | -4.3658
2.3868 | -34.0264 | -4.3753
2.3929 | -34.0268 | -4.3852
2.399 | -34.027 | -4.3939
2.4051 | -34.0275 | -4.4039
2.4112 | -34.0279 | -4.4133
2.4173 | -34.0282 | -4.4226
2.4234 | -34.0282 | -4.4314
2.4295 | -34.0285 | -4.4406
2.4356 | -34.0287 | -4.4504
2.4417 | -34.0287 | -4.4608
2.4478 | -34.0288 | -4.4708
2.4539 | -34.0291 | -4.479
2.46 | -34.0288 | -4.4889
2.4661 | -34.0288 | -4.4985
2.4722 | -34.0291 | -4.5075
2.4783 | -34.0292 | -4.5162
2.4844 | -34.0292 | -4.5252
2.4905 | -34.0291 | -4.5346
2.4966 | -34.0289 | -4.5433
2.5027 | -34.0289 | -4.5516
2.5089 | -34.0288 | -4.5602
2.515 | -34.0286 | -4.5696
2.5211 | -34.0286 | -4.5779
2.5272 | -34.0282 | -4.5867
2.5333 | -34.0281 | -4.5964
2.5394 | -34.028 | -4.6057
2.5455 | -34.0279 | -4.6152
2.5516 | -34.0278 | -4.6239
2.5577 | -34.0278 | -4.634
2.5638 | -34.0275 | -4.644
2.5699 | -34.0273 | -4.6543
2.576 | -34.027 | -4.6648
2.5821 | -34.0269 | -4.6745
2.5882 | -34.0267 | -4.684
2.5943 | -34.0265 | -4.6935
2.6004 | -34.0263 | -4.7028
2.6065 | -34.0261 | -4.7111
2.6126 | -34.026 | -4.7202
2.6187 | -34.0258 | -4.7288
2.6248 | -34.0257 | -4.7377
2.6309 | -34.0255 | -4.7486
2.637 | -34.0254 | -4.7593
2.6431 | -34.0254 | -4.7686
2.6492 | -34.0253 | -4.7782
2.6554 | -34.0253 | -4.7882
2.6615 | -34.025 | -4.7975
2.6676 | -34.025 | -4.8078
2.6737 | -34.0251 | -4.8182
2.6798 | -34.0251 | -4.8277
2.6859 | -34.0251 | -4.8378
2.692 | -34.0249 | -4.8473
2.6981 | -34.0249 | -4.8581
2.7042 | -34.0249 | -4.8697
2.7103 | -34.0249 | -4.8794
2.7164 | -34.0251 | -4.8898
2.7225 | -34.0252 | -4.9008
2.7286 | -34.0252 | -4.9116
2.7347 | -34.0252 | -4.9212
2.7408 | -34.0255 | -4.93
2.7469 | -34.0254 | -4.9411
2.753 | -34.0257 | -4.9514
2.7591 | -34.0258 | -4.9631
2.7652 | -34.0258 | -4.9743
2.7713 | -34.0258 | -4.9838
2.7774 | -34.0261 | -4.995
2.7835 | -34.0262 | -5.0054
2.7896 | -34.0263 | -5.0144
2.7958 | -34.0267 | -5.0244
2.8019 | -34.0272 | -5.0339
2.808 | -34.0275 | -5.0456
2.8141 | -34.0279 | -5.0584
2.8202 | -34.0283 | -5.0703
2.8263 | -34.0287 | -5.0809
2.8324 | -34.0289 | -5.0931
2.8385 | -34.0293 | -5.1051
2.8446 | -34.0299 | -5.1166
2.8507 | -34.0302 | -5.1285
2.8568 | -34.0307 | -5.1395
2.8629 | -34.0311 | -5.1505
2.869 | -34.0315 | -5.1617
2.8751 | -34.0319 | -5.1735
2.8812 | -34.0326 | -5.1868
2.8873 | -34.0329 | -5.1981
2.8934 | -34.0333 | -5.2099
2.8995 | -34.0337 | -5.2218
2.9056 | -34.034 | -5.2325
2.9117 | -34.0345 | -5.2439
2.9178 | -34.0352 | -5.2567
2.9239 | -34.0356 | -5.2697
2.93 | -34.0361 | -5.283
2.9361 | -34.0367 | -5.2942
2.9423 | -34.0373 | -5.3061
2.9484 | -34.0378 | -5.3196
2.9545 | -34.0386 | -5.3328
2.9606 | -34.0391 | -5.3461
2.9667 | -34.0396 | -5.3578
2.9728 | -34.04 | -5.3699
2.9789 | -34.0405 | -5.3825
2.985 | -34.041 | -5.3963
2.9911 | -34.0415 | -5.4087
2.9972 | -34.0421 | -5.4213
3.0033 | -34.0427 | -5.4341
3.0094 | -34.0433 | -5.447
3.0155 | -34.0439 | -5.4595
3.0216 | -34.0444 | -5.4729
3.0277 | -34.0449 | -5.4865
3.0338 | -34.0456 | -5.5
3.0399 | -34.0462 | -5.5136
3.046 | -34.0465 | -5.5266
3.0521 | -34.0472 | -5.5391
3.0582 | -34.0475 | -5.553
3.0643 | -34.0481 | -5.5662
3.0704 | -34.0484 | -5.5795
3.0765 | -34.0488 | -5.5917
3.0827 | -34.0492 | -5.6048
3.0888 | -34.0496 | -5.6175
3.0949 | -34.05 | -5.6305
3.101 | -34.0505 | -5.6445
3.1071 | -34.0509 | -5.6588
3.1132 | -34.0514 | -5.6727
3.1193 | -34.0519 | -5.6858
3.1254 | -34.0523 | -5.698
3.1315 | -34.0528 | -5.7122
3.1376 | -34.0531 | -5.7251
3.1437 | -34.0534 | -5.7387
3.1498 | -34.0539 | -5.7529
3.1559 | -34.0542 | -5.7661
3.162 | -34.0543 | -5.7803
3.1681 | -34.0544 | -5.7946
3.1742 | -34.0549 | -5.8082
3.1803 | -34.0553 | -5.8215
3.1864 | -34.0557 | -5.8337
3.1925 | -34.0561 | -5.8475
3.1986 | -34.0562 | -5.8602
3.2047 | -34.0564 | -5.8741
3.2108 | -34.0566 | -5.8859
3.2169 | -34.0571 | -5.8996
3.223 | -34.0571 | -5.9128
3.2292 | -34.0571 | -5.9267
3.2353 | -34.0571 | -5.9362
3.2414 | -34.0573 | -5.9492
3.2475 | -34.0573 | -5.9601
3.2536 | -34.0577 | -5.9739
3.2597 | -34.0579 | -5.985
3.2658 | -34.0581 | -5.9957
3.2719 | -34.0582 | -6.0067
3.278 | -34.0584 | -6.0188
3.2841 | -34.0585 | -6.0307
3.2902 | -34.0586 | -6.0434
3.2963 | -34.0588 | -6.0549
3.3024 | -34.0587 | -6.065
3.3085 | -34.059 | -6.0774
3.3146 | -34.0589 | -6.0884
3.3207 | -34.0591 | -6.1005
3.3268 | -34.0592 | -6.1126
3.3329 | -34.0594 | -6.1225
3.339 | -34.0595 | -6.1321
3.3451 | -34.0595 | -6.1436
3.3512 | -34.0597 | -6.1531
3.3573 | -34.0598 | -6.1644
3.3634 | -34.0599 | -6.1759
3.3696 | -34.0599 | -6.1878
3.3757 | -34.0599 | -6.1977
3.3818 | -34.0602 | -6.2094
3.3879 | -34.0602 | -6.2196
3.394 | -34.0601 | -6.2302
3.4001 | -34.0602 | -6.2409
3.4062 | -34.0602 | -6.2512
3.4123 | -34.0602 | -6.2609
3.4184 | -34.0602 | -6.2719
3.4245 | -34.0603 | -6.2824
3.4306 | -34.0603 | -6.2921
3.4367 | -34.0604 | -6.3019
3.4428 | -34.0607 | -6.3111
3.4489 | -34.0606 | -6.3224
3.455 | -34.0605 | -6.3329
3.4611 | -34.0605 | -6.344
3.4672 | -34.0605 | -6.3552
3.4733 | -34.0606 | -6.3654
3.4794 | -34.0606 | -6.3759
3.4855 | -34.0608 | -6.3856
3.4916 | -34.0609 | -6.3953
3.4977 | -34.0608 | -6.4062
3.5038 | -34.061 | -6.4161
3.5099 | -34.061 | -6.4257
3.5161 | -34.0611 | -6.4353
3.5222 | -34.0612 | -6.4452
3.5283 | -34.0611 | -6.4553
3.5344 | -34.0611 | -6.466
3.5405 | -34.0613 | -6.476
3.5466 | -34.0614 | -6.4858
3.5527 | -34.0614 | -6.4965
3.5588 | -34.0615 | -6.5064
3.5649 | -34.0615 | -6.5166
3.571 | -34.0616 | -6.5269
3.5771 | -34.0617 | -6.5377
3.5832 | -34.0617 | -6.5474
3.5893 | -34.062 | -6.5573
3.5954 | -34.062 | -6.5679
3.6015 | -34.0621 | -6.5783
3.6076 | -34.0624 | -6.588
3.6137 | -34.0624 | -6.598
3.6198 | -34.0627 | -6.6072
3.6259 | -34.0628 | -6.6179
3.632 | -34.063 | -6.6287
3.6381 | -34.0632 | -6.6386
3.6442 | -34.0634 | -6.649
3.6503 | -34.0636 | -6.6582
3.6565 | -34.0639 | -6.6701
3.6626 | -34.0639 | -6.6808
3.6687 | -34.0642 | -6.6916
3.6748 | -34.0645 | -6.7016
3.6809 | -34.0649 | -6.7122
3.687 | -34.065 | -6.7236
3.6931 | -34.0654 | -6.7337
3.6992 | -34.0656 | -6.7444
3.7053 | -34.0658 | -6.7547
3.7114 | -34.0659 | -6.7661
3.7175 | -34.0663 | -6.7761
3.7236 | -34.0665 | -6.7852
3.7297 | -34.0666 | -6.7959
3.7358 | -34.0668 | -6.8078
3.7419 | -34.067 | -6.8195
3.748 | -34.0672 | -6.8305
3.7541 | -34.0673 | -6.8417
3.7602 | -34.0676 | -6.8527
3.7663 | -34.0678 | -6.8631
3.7724 | -34.0681 | -6.8741
3.7785 | -34.0684 | -6.8846
3.7846 | -34.0686 | -6.8954
3.7907 | -34.069 | -6.9066
3.7969 | -34.0692 | -6.9183
3.803 | -34.0696 | -6.9285
3.8091 | -34.0698 | -6.939
3.8152 | -34.0701 | -6.9503
3.8213 | -34.0703 | -6.9622
3.8274 | -34.0706 | -6.9728
3.8335 | -34.071 | -6.9846
3.8396 | -34.0715 | -6.9963
3.8457 | -34.0718 | -7.0067
3.8518 | -34.0723 | -7.0177
3.8579 | -34.0726 | -7.0285
3.864 | -34.073 | -7.0394
3.8701 | -34.0734 | -7.0505
3.8762 | -34.0737 | -7.0618
3.8823 | -34.0739 | -7.0726
3.8884 | -34.0744 | -7.0832
3.8945 | -34.0748 | -7.0948
3.9006 | -34.0752 | -7.1064
3.9067 | -34.0755 | -7.1185
3.9128 | -34.0759 | -7.13
3.9189 | -34.0763 | -7.1408
3.925 | -34.0765 | -7.1518
3.9311 | -34.0769 | -7.162
3.9372 | -34.0773 | -7.1737
3.9434 | -34.0777 | -7.1844
3.9495 | -34.078 | -7.1962
3.9556 | -34.0785 | -7.2075
3.9617 | -34.0788 | -7.2185
3.9678 | -34.0792 | -7.2294
3.9739 | -34.0796 | -7.2407
3.98 | -34.08 | -7.2528
3.9861 | -34.0802 | -7.2633
3.9922 | -34.0805 | -7.2752
3.9983 | -34.0809 | -7.2854
4.0044 | -34.0812 | -7.2953
4.0105 | -34.0815 | -7.3063
4.0166 | -34.082 | -7.3175
4.0227 | -34.0824 | -7.328
4.0288 | -34.0826 | -7.3396
4.0349 | -34.0829 | -7.3496
4.041 | -34.0833 | -7.3606
4.0471 | -34.0835 | -7.3711
4.0532 | -34.084 | -7.3824
4.0593 | -34.0842 | -7.394
4.0654 | -34.0846 | -7.4049
4.0715 | -34.0849 | -7.4164
4.0776 | -34.0852 | -7.4268
4.0838 | -34.0856 | -7.4384
4.0899 | -34.0859 | -7.4497
4.096 | -34.0862 | -7.4603
4.1021 | -34.0865 | -7.4711
4.1082 | -34.0869 | -7.4815
4.1143 | -34.0872 | -7.4915
4.1204 | -34.0877 | -7.5023
4.1265 | -34.088 | -7.5128
4.1326 | -34.0884 | -7.5232
4.1387 | -34.0888 | -7.5344
4.1448 | -34.0891 | -7.5455
4.1509 | -34.0894 | -7.5563
4.157 | -34.0896 | -7.5675
4.1631 | -34.0899 | -7.5788
4.1692 | -34.0903 | -7.5896
4.1753 | -34.0907 | -7.6005
4.1814 | -34.0909 | -7.6103
4.1875 | -34.0912 | -7.6218
4.1936 | -34.0915 | -7.6333
4.1997 | -34.0918 | -7.6434
4.2058 | -34.0921 | -7.655
4.2119 | -34.0924 | -7.6648
4.218 | -34.0926 | -7.6749
4.2241 | -34.093 | -7.6845
4.2303 | -34.0934 | -7.6958
4.2364 | -34.0937 | -7.7072
4.2425 | -34.094 | -7.7177
4.2486 | -34.0942 | -7.729
4.2547 | -34.0944 | -7.7382
4.2608 | -34.0946 | -7.7481
4.2669 | -34.0949 | -7.7585
4.273 | -34.0952 | -7.7692
4.2791 | -34.0954 | -7.78
4.2852 | -34.0956 | -7.7903
4.2913 | -34.0958 | -7.7994
4.2974 | -34.096 | -7.8103
4.3035 | -34.0962 | -7.8211
4.3096 | -34.0964 | -7.8315
4.3157 | -34.0966 | -7.8414
4.3218 | -34.0968 | -7.8513
4.3279 | -34.0971 | -7.8616
4.334 | -34.0973 | -7.8714
4.3401 | -34.0975 | -7.8813
4.3462 | -34.0978 | -7.8921
4.3523 | -34.0981 | -7.9017
4.3584 | -34.0982 | -7.9101
4.3645 | -34.0984 | -7.9197
4.3707 | -34.0987 | -7.9289
4.3768 | -34.0988 | -7.9392
4.3829 | -34.099 | -7.9495
4.389 | -34.0993 | -7.9591
4.3951 | -34.0995 | -7.9685
4.4012 | -34.0997 | -7.9784
4.4073 | -34.0999 | -7.9884
4.4134 | -34.1001 | -7.9985
4.4195 | -34.1004 | -8.0081
4.4256 | -34.1007 | -8.0175
4.4317 | -34.1009 | -8.0269
4.4378 | -34.101 | -8.0364
4.4439 | -34.1012 | -8.0457
4.45 | -34.1013 | -8.0556
4.4561 | -34.1014 | -8.0653
4.4622 | -34.1016 | -8.0748
4.4683 | -34.1017 | -8.084
4.4744 | -34.1019 | -8.0941
4.4805 | -34.1022 | -8.1032
4.4866 | -34.1024 | -8.1123
4.4927 | -34.1026 | -8.1224
4.4988 | -34.1028 | -8.1317
4.5049 | -34.103 | -8.1407
4.511 | -34.1033 | -8.1496
4.5172 | -34.1036 | -8.1593
4.5233 | -34.1037 | -8.1698
4.5294 | -34.104 | -8.179
4.5355 | -34.1042 | -8.1883
4.5416 | -34.1044 | -8.1974
4.5477 | -34.1046 | -8.2065
4.5538 | -34.1048 | -8.2152
4.5599 | -34.1051 | -8.2246
4.566 | -34.1053 | -8.2343
4.5721 | -34.1055 | -8.2434
4.5782 | -34.1058 | -8.2531
4.5843 | -34.106 | -8.2626
4.5904 | -34.1063 | -8.2718
4.5965 | -34.1066 | -8.2804
4.6026 | -34.1069 | -8.291
4.6087 | -34.107 | -8.2999
4.6148 | -34.1073 | -8.3105
4.6209 | -34.1076 | -8.3188
4.627 | -34.108 | -8.3288
4.6331 | -34.108 | -8.3376
4.6392 | -34.1083 | -8.3467
4.6453 | -34.1085 | -8.356
4.6514 | -34.1088 | -8.3655
4.6576 | -34.1088 | -8.3759
4.6637 | -34.1091 | -8.3835
4.6698 | -34.1093 | -8.392
4.6759 | -34.1094 | -8.4016
4.682 | -34.1096 | -8.4107
4.6881 | -34.1098 | -8.4194
4.6942 | -34.11 | -8.4286
4.7003 | -34.1103 | -8.438
4.7064 | -34.1105 | -8.4472
4.7125 | -34.1108 | -8.4567
4.7186 | -34.1109 | -8.465
4.7247 | -34.1113 | -8.4746
4.7308 | -34.1116 | -8.4838
4.7369 | -34.1118 | -8.4933
4.743 | -34.112 | -8.5036
4.7491 | -34.1122 | -8.5132
4.7552 | -34.1123 | -8.5223
4.7613 | -34.1125 | -8.5319
4.7674 | -34.1128 | -8.5404
4.7735 | -34.1129 | -8.5504
4.7796 | -34.1131 | -8.5602
4.7857 | -34.1132 | -8.5698
4.7918 | -34.1135 | -8.5783
4.7979 | -34.1138 | -8.5873
4.8041 | -34.1139 | -8.5966
4.8102 | -34.1142 | -8.6056
4.8163 | -34.1145 | -8.6151
4.8224 | -34.1146 | -8.6237
4.8285 | -34.1149 | -8.6326
4.8346 | -34.1151 | -8.6421
4.8407 | -34.1153 | -8.6515
4.8468 | -34.1156 | -8.6612
4.8529 | -34.1158 | -8.671
4.859 | -34.116 | -8.6804
4.8651 | -34.1162 | -8.6898
4.8712 | -34.1165 | -8.6991
4.8773 | -34.1167 | -8.7086
4.8834 | -34.1169 | -8.7183
4.8895 | -34.1172 | -8.7284
4.8956 | -34.1174 | -8.7383
4.9017 | -34.1175 | -8.7477
4.9078 | -34.1177 | -8.7577
4.9139 | -34.118 | -8.7661
4.92 | -34.1182 | -8.7752
4.9261 | -34.1186 | -8.784
4.9322 | -34.1189 | -8.7931
4.9383 | -34.1192 | -8.8022
4.9445 | -34.1195 | -8.8123
4.9506 | -34.1197 | -8.8211
4.9567 | -34.12 | -8.8305
4.9628 | -34.1202 | -8.8399
4.9689 | -34.1204 | -8.8497
4.975 | -34.1206 | -8.859
4.9811 | -34.1209 | -8.8689
4.9872 | -34.1211 | -8.8787
4.9933 | -34.1215 | -8.8885
4.9994 | -34.1218 | -8.8984
5.0055 | -34.1222 | -8.908
5.0116 | -34.1225 | -8.9183
5.0177 | -34.1229 | -8.9284
5.0238 | -34.1232 | -8.9389
5.0299 | -34.1236 | -8.9499
5.036 | -34.1239 | -8.9595
5.0421 | -34.1241 | -8.9692
5.0482 | -34.1244 | -8.9795
5.0543 | -34.1247 | -8.9889
5.0604 | -34.125 | -8.9992
5.0665 | -34.1254 | -9.0089
5.0726 | -34.1257 | -9.0193
5.0787 | -34.126 | -9.0295
5.0848 | -34.1263 | -9.0396
5.091 | -34.1266 | -9.0496
5.0971 | -34.1269 | -9.0597
5.1032 | -34.1273 | -9.0701
5.1093 | -34.1276 | -9.0801
5.1154 | -34.1278 | -9.0901
5.1215 | -34.1281 | -9.1001
5.1276 | -34.1284 | -9.11
5.1337 | -34.1289 | -9.1197
5.1398 | -34.1291 | -9.1294
5.1459 | -34.1295 | -9.1399
5.152 | -34.1298 | -9.151
5.1581 | -34.13 | -9.1617
5.1642 | -34.1303 | -9.1716
5.1703 | -34.1307 | -9.1817
5.1764 | -34.131 | -9.1916
5.1825 | -34.1312 | -9.202
5.1886 | -34.1315 | -9.2119
5.1947 | -34.1318 | -9.2227
5.2008 | -34.1323 | -9.2326
5.2069 | -34.1326 | -9.2437
5.213 | -34.1329 | -9.2534
5.2191 | -34.1333 | -9.2638
5.2252 | -34.1336 | -9.2741
5.2314 | -34.1338 | -9.2849
5.2375 | -34.1342 | -9.2949
5.2436 | -34.1345 | -9.3051
5.2497 | -34.1348 | -9.315
5.2558 | -34.1352 | -9.326
5.2619 | -34.1355 | -9.337
5.268 | -34.1358 | -9.3474
5.2741 | -34.1362 | -9.3586
5.2802 | -34.1366 | -9.3695
5.2863 | -34.1369 | -9.3798
5.2924 | -34.1372 | -9.391
5.2985 | -34.1377 | -9.4023
5.3046 | -34.1381 | -9.4136
5.3107 | -34.1385 | -9.4245
5.3168 | -34.1389 | -9.4349
5.3229 | -34.1392 | -9.4448
5.329 | -34.1396 | -9.4548
5.3351 | -34.1401 | -9.465
5.3412 | -34.1405 | -9.4751
5.3473 | -34.1408 | -9.486
5.3534 | -34.1412 | -9.4969
5.3595 | -34.1415 | -9.5077
5.3656 | -34.1419 | -9.5187
5.3717 | -34.1422 | -9.5293
5.3779 | -34.1426 | -9.5402
5.384 | -34.1429 | -9.551
5.3901 | -34.1432 | -9.5617
5.3962 | -34.1435 | -9.5725
5.4023 | -34.1439 | -9.5833
5.4084 | -34.1443 | -9.5938
5.4145 | -34.1446 | -9.6047
5.4206 | -34.1449 | -9.616
5.4267 | -34.1453 | -9.6273
5.4328 | -34.1457 | -9.6375
5.4389 | -34.1461 | -9.6478
5.445 | -34.1463 | -9.6591
5.4511 | -34.1467 | -9.6695
5.4572 | -34.1472 | -9.6807
5.4633 | -34.1475 | -9.6911
5.4694 | -34.1477 | -9.7017
5.4755 | -34.1481 | -9.7125
5.4816 | -34.1484 | -9.7231
5.4877 | -34.1488 | -9.7338
5.4938 | -34.1491 | -9.7442
5.4999 | -34.1495 | -9.7549
5.506 | -34.1499 | -9.7657
5.5121 | -34.1502 | -9.7768
5.5183 | -34.1505 | -9.7868
5.5244 | -34.1508 | -9.7976
5.5305 | -34.1512 | -9.8085
5.5366 | -34.1515 | -9.8189
5.5427 | -34.1518 | -9.8302
5.5488 | -34.1522 | -9.8412
5.5549 | -34.1524 | -9.8519
5.561 | -34.1528 | -9.8632
5.5671 | -34.1532 | -9.8736
5.5732 | -34.1535 | -9.8845
5.5793 | -34.1539 | -9.896
5.5854 | -34.1543 | -9.9069
5.5915 | -34.1545 | -9.9178
5.5976 | -34.1549 | -9.9288
5.6037 | -34.1552 | -9.9395
5.6098 | -34.1555 | -9.95
5.6159 | -34.1559 | -9.9609
5.622 | -34.1562 | -9.9714
5.6281 | -34.1566 | -9.9822
5.6342 | -34.1569 | -9.9931
5.6403 | -34.1572 | -10.0034
5.6464 | -34.1576 | -10.0142
5.6525 | -34.1579 | -10.0247
5.6586 | -34.1583 | -10.0346
5.6648 | -34.1586 | -10.0459
5.6709 | -34.1589 | -10.057
5.677 | -34.1593 | -10.0684
5.6831 | -34.1596 | -10.0786
5.6892 | -34.1599 | -10.0894
5.6953 | -34.1602 | -10.0999
5.7014 | -34.1606 | -10.1112
5.7075 | -34.1609 | -10.1221
5.7136 | -34.1612 | -10.1326
5.7197 | -34.1615 | -10.1432
5.7258 | -34.1618 | -10.1537
5.7319 | -34.1621 | -10.1644
5.738 | -34.1625 | -10.1751
5.7441 | -34.1628 | -10.1858
5.7502 | -34.1632 | -10.196
5.7563 | -34.1636 | -10.2065
5.7624 | -34.1639 | -10.2179
5.7685 | -34.1641 | -10.2285
5.7746 | -34.1644 | -10.239
5.7807 | -34.1647 | -10.2493
5.7868 | -34.165 | -10.26
5.7929 | -34.1653 | -10.2705
5.799 | -34.1656 | -10.281
5.8052 | -34.1659 | -10.2918
5.8113 | -34.1662 | -10.3024
5.8174 | -34.1664 | -10.3133
5.8235 | -34.1667 | -10.3237
5.8296 | -34.167 | -10.3349
5.8357 | -34.1672 | -10.3455
5.8418 | -34.1676 | -10.3564
5.8479 | -34.1678 | -10.3673
5.854 | -34.1681 | -10.3782
5.8601 | -34.1684 | -10.3886
5.8662 | -34.1687 | -10.3996
5.8723 | -34.169 | -10.4095
5.8784 | -34.1693 | -10.4201
5.8845 | -34.1696 | -10.4309
5.8906 | -34.1699 | -10.4414
5.8967 | -34.1702 | -10.4516
5.9028 | -34.1705 | -10.4622
5.9089 | -34.1708 | -10.4722
5.915 | -34.1711 | -10.4828
5.9211 | -34.1714 | -10.4934
5.9272 | -34.1717 | -10.5039
5.9333 | -34.172 | -10.5137
5.9394 | -34.1724 | -10.5241
5.9455 | -34.1727 | -10.5343
5.9517 | -34.173 | -10.5448
5.9578 | -34.1734 | -10.5557
5.9639 | -34.1737 | -10.5658
5.97 | -34.174 | -10.5765
5.9761 | -34.1743 | -10.5867
5.9822 | -34.1747 | -10.5974
5.9883 | -34.175 | -10.6077
5.9944 | -34.1753 | -10.6182
6.0005 | -34.1757 | -10.629
6.0066 | -34.176 | -10.6391
6.0127 | -34.1763 | -10.6489
6.0188 | -34.1766 | -10.6598
6.0249 | -34.1769 | -10.6702
6.031 | -34.1772 | -10.6804
6.0371 | -34.1776 | -10.6911
6.0432 | -34.1779 | -10.7015
6.0493 | -34.1783 | -10.7118
6.0554 | -34.1787 | -10.7221
6.0615 | -34.179 | -10.7331
6.0676 | -34.1793 | -10.7442
6.0737 | -34.1796 | -10.7548
6.0798 | -34.1799 | -10.7651
6.0859 | -34.1802 | -10.7756
6.0921 | -34.1806 | -10.7857
6.0982 | -34.1809 | -10.7964
6.1043 | -34.1813 | -10.8067
6.1104 | -34.1817 | -10.8177
6.1165 | -34.182 | -10.8273
6.1226 | -34.1823 | -10.8372
6.1287 | -34.1827 | -10.8477
6.1348 | -34.1831 | -10.8581
6.1409 | -34.1834 | -10.8684
6.147 | -34.1838 | -10.8787
6.1531 | -34.1841 | -10.8891
6.1592 | -34.1844 | -10.8995
6.1653 | -34.1847 | -10.9097
6.1714 | -34.1851 | -10.9201
6.1775 | -34.1854 | -10.9308
6.1836 | -34.1857 | -10.941
6.1897 | -34.1861 | -10.9515
6.1958 | -34.1863 | -10.9621
6.2019 | -34.1867 | -10.9728
6.208 | -34.1871 | -10.9835
6.2141 | -34.1874 | -10.9942
6.2202 | -34.1878 | -11.005
6.2263 | -34.1881 | -11.0154
6.2325 | -34.1885 | -11.0253
6.2386 | -34.1889 | -11.0353
6.2447 | -34.1892 | -11.046
6.2508 | -34.1895 | -11.0563
6.2569 | -34.1899 | -11.0671
6.263 | -34.1902 | -11.0771
6.2691 | -34.1906 | -11.0874
6.2752 | -34.191 | -11.0977
6.2813 | -34.1913 | -11.1081
6.2874 | -34.1916 | -11.1184
6.2935 | -34.192 | -11.1286
6.2996 | -34.1923 | -11.1391
6.3057 | -34.1927 | -11.1494
6.3118 | -34.193 | -11.1597
6.3179 | -34.1934 | -11.1701
6.324 | -34.1937 | -11.1807
6.3301 | -34.1941 | -11.1909
6.3362 | -34.1945 | -11.2013
6.3423 | -34.1949 | -11.2118
6.3484 | -34.1953 | -11.2224
6.3545 | -34.1957 | -11.2321
6.3606 | -34.1961 | -11.2422
6.3667 | -34.1964 | -11.2524
6.3728 | -34.1968 | -11.2628
6.379 | -34.1971 | -11.2731
6.3851 | -34.1975 | -11.2836
6.3912 | -34.1979 | -11.2937
6.3973 | -34.1983 | -11.3038
6.4034 | -34.1986 | -11.314
6.4095 | -34.1989 | -11.3246
6.4156 | -34.1993 | -11.3346
6.4217 | -34.1996 | -11.345
6.4278 | -34.2 | -11.3552
6.4339 | -34.2004 | -11.3656
6.44 | -34.2008 | -11.3755
6.4461 | -34.2011 | -11.3863
6.4522 | -34.2015 | -11.396
6.4583 | -34.2018 | -11.4064
6.4644 | -34.202 | -11.4171
6.4705 | -34.2025 | -11.4275
6.4766 | -34.2028 | -11.4382
6.4827 | -34.2032 | -11.449
6.4888 | -34.2035 | -11.4592
6.4949 | -34.2039 | -11.469
6.501 | -34.2042 | -11.4797
6.5071 | -34.2046 | -11.4902
6.5132 | -34.205 | -11.5004
6.5194 | -34.2054 | -11.5107
6.5255 | -34.2058 | -11.5206
6.5316 | -34.2061 | -11.5312
6.5377 | -34.2064 | -11.5417
6.5438 | -34.2068 | -11.552
6.5499 | -34.2072 | -11.5621
6.556 | -34.2075 | -11.5722
6.5621 | -34.2078 | -11.5827
6.5682 | -34.2082 | -11.5929
6.5743 | -34.2085 | -11.604
6.5804 | -34.2088 | -11.6144
6.5865 | -34.2091 | -11.6249
6.5926 | -34.2095 | -11.6352
6.5987 | -34.2099 | -11.6451
6.6048 | -34.2102 | -11.6561
6.6109 | -34.2106 | -11.6664
6.617 | -34.2109 | -11.6775
6.6231 | -34.2112 | -11.6881
6.6292 | -34.2117 | -11.6983
6.6353 | -34.212 | -11.7096
6.6414 | -34.2124 | -11.7203
6.6475 | -34.2128 | -11.7308
6.6536 | -34.2131 | -11.7408
6.6597 | -34.2136 | -11.751
6.6659 | -34.2139 | -11.7611
6.672 | -34.2143 | -11.7715
6.6781 | -34.2146 | -11.7816
6.6842 | -34.215 | -11.7922
6.6903 | -34.2154 | -11.8024
6.6964 | -34.2158 | -11.8129
6.7025 | -34.2161 | -11.8232
6.7086 | -34.2164 | -11.8333
6.7147 | -34.2168 | -11.8438
6.7208 | -34.2171 | -11.8546
6.7269 | -34.2175 | -11.8646
6.733 | -34.2177 | -11.8742
6.7391 | -34.2182 | -11.8847
6.7452 | -34.2186 | -11.8945
6.7513 | -34.219 | -11.9045
6.7574 | -34.2194 | -11.9151
6.7635 | -34.2198 | -11.9249
6.7696 | -34.2202 | -11.9353
6.7757 | -34.2206 | -11.9459
6.7818 | -34.2209 | -11.9563
6.7879 | -34.2213 | -11.9667
6.794 | -34.2217 | -11.9768
6.8001 | -34.222 | -11.987
6.8063 | -34.2224 | -11.9975
6.8124 | -34.2227 | -12.0079
6.8185 | -34.223 | -12.0183
6.8246 | -34.2234 | -12.0286
6.8307 | -34.2237 | -12.0387
6.8368 | -34.2241 | -12.0493
6.8429 | -34.2245 | -12.0595
6.849 | -34.2248 | -12.0698
6.8551 | -34.2252 | -12.0801
6.8612 | -34.2256 | -12.0902
6.8673 | -34.226 | -12.1002
6.8734 | -34.2263 | -12.1113
6.8795 | -34.2267 | -12.1218
6.8856 | -34.2271 | -12.1315
6.8917 | -34.2275 | -12.1412
6.8978 | -34.2278 | -12.1514
6.9039 | -34.2281 | -12.1615
6.91 | -34.2285 | -12.1719
6.9161 | -34.2289 | -12.182
6.9222 | -34.2292 | -12.1923
6.9283 | -34.2295 | -12.2022
6.9344 | -34.2299 | -12.2123
6.9405 | -34.2302 | -12.2223
6.9466 | -34.2306 | -12.2327
6.9528 | -34.2309 | -12.243
6.9589 | -34.2313 | -12.2531
6.965 | -34.2317 | -12.2633
6.9711 | -34.232 | -12.2735
6.9772 | -34.2325 | -12.2835
6.9833 | -34.2327 | -12.2939
6.9894 | -34.2331 | -12.3045
6.9955 | -34.2335 | -12.3146
7.0016 | -34.2338 | -12.3252
7.0077 | -34.2342 | -12.3357
7.0138 | -34.2346 | -12.3462
7.0199 | -34.2349 | -12.3562
7.026 | -34.2351 | -12.3662
7.0321 | -34.2355 | -12.3764
7.0382 | -34.2358 | -12.3863
7.0443 | -34.2362 | -12.3968
7.0504 | -34.2365 | -12.4073
7.0565 | -34.2369 | -12.4177
7.0626 | -34.2372 | -12.4276
7.0687 | -34.2375 | -12.4374
7.0748 | -34.2379 | -12.4474
7.0809 | -34.2382 | -12.4578
7.087 | -34.2386 | -12.468
7.0932 | -34.2389 | -12.4782
7.0993 | -34.2392 | -12.4884
7.1054 | -34.2396 | -12.4982
7.1115 | -34.2399 | -12.5082
7.1176 | -34.2402 | -12.5182
7.1237 | -34.2404 | -12.5283
7.1298 | -34.2407 | -12.538
7.1359 | -34.241 | -12.5478
7.142 | -34.2413 | -12.5579
7.1481 | -34.2417 | -12.5677
7.1542 | -34.242 | -12.5777
7.1603 | -34.2423 | -12.5877
7.1664 | -34.2427 | -12.5979
7.1725 | -34.243 | -12.6077
7.1786 | -34.2434 | -12.6178
7.1847 | -34.2437 | -12.6277
7.1908 | -34.244 | -12.6379
7.1969 | -34.2443 | -12.6479
7.203 | -34.2447 | -12.658
7.2091 | -34.245 | -12.6678
7.2152 | -34.2453 | -12.6777
7.2213 | -34.2456 | -12.6875
7.2274 | -34.2459 | -12.6975
7.2335 | -34.2463 | -12.7075
7.2397 | -34.2466 | -12.7175
7.2458 | -34.2469 | -12.7271
7.2519 | -34.2472 | -12.7371
7.258 | -34.2476 | -12.7473
7.2641 | -34.2479 | -12.7572
7.2702 | -34.2482 | -12.7674
7.2763 | -34.2486 | -12.7772
7.2824 | -34.249 | -12.7871
7.2885 | -34.2493 | -12.797
7.2946 | -34.2496 | -12.8071
7.3007 | -34.2499 | -12.8172
7.3068 | -34.2502 | -12.8274
7.3129 | -34.2505 | -12.8371
7.319 | -34.2509 | -12.8472
7.3251 | -34.2512 | -12.8568
7.3312 | -34.2515 | -12.8664
7.3373 | -34.2519 | -12.8761
7.3434 | -34.2523 | -12.8858
7.3495 | -34.2526 | -12.8952
7.3556 | -34.2529 | -12.9055
7.3617 | -34.2532 | -12.9159
7.3678 | -34.2536 | -12.9259
7.3739 | -34.2539 | -12.9354
7.3801 | -34.2542 | -12.9452
7.3862 | -34.2545 | -12.955
7.3923 | -34.2549 | -12.9648
7.3984 | -34.2552 | -12.9746
7.4045 | -34.2555 | -12.9844
7.4106 | -34.2559 | -12.9941
7.4167 | -34.2561 | -13.0042
7.4228 | -34.2565 | -13.0139
7.4289 | -34.2568 | -13.0235
7.435 | -34.2572 | -13.0331
7.4411 | -34.2575 | -13.0426
7.4472 | -34.2579 | -13.0526
7.4533 | -34.2582 | -13.0624
7.4594 | -34.2585 | -13.0725
7.4655 | -34.2589 | -13.0823
7.4716 | -34.2592 | -13.0924
7.4777 | -34.2595 | -13.1017
7.4838 | -34.2599 | -13.1117
7.4899 | -34.2603 | -13.1215
7.496 | -34.2606 | -13.1314
7.5021 | -34.2609 | -13.141
7.5082 | -34.2612 | -13.1509
7.5143 | -34.2616 | -13.1606
7.5204 | -34.2619 | -13.1701
7.5266 | -34.2622 | -13.1799
7.5327 | -34.2625 | -13.1896
7.5388 | -34.2628 | -13.1993
7.5449 | -34.2631 | -13.2094
7.551 | -34.2634 | -13.2195
7.5571 | -34.2638 | -13.2291
7.5632 | -34.2641 | -13.2388
7.5693 | -34.2645 | -13.2488
7.5754 | -34.2648 | -13.2585
7.5815 | -34.2652 | -13.268
7.5876 | -34.2654 | -13.2778
7.5937 | -34.2658 | -13.2874
7.5998 | -34.2661 | -13.2973
7.6059 | -34.2664 | -13.307
7.612 | -34.2667 | -13.3167
7.6181 | -34.267 | -13.3264
7.6242 | -34.2674 | -13.3362
7.6303 | -34.2678 | -13.3459
7.6364 | -34.268 | -13.3559
7.6425 | -34.2684 | -13.366
7.6486 | -34.2687 | -13.3759
7.6547 | -34.2691 | -13.3857
7.6608 | -34.2694 | -13.3955
7.667 | -34.2697 | -13.4047
7.6731 | -34.27 | -13.4146
7.6792 | -34.2703 | -13.4245
7.6853 | -34.2706 | -13.4342
7.6914 | -34.2709 | -13.4441
7.6975 | -34.2712 | -13.454
7.7036 | -34.2715 | -13.4639
7.7097 | -34.2718 | -13.4736
7.7158 | -34.2721 | -13.4832
7.7219 | -34.2724 | -13.4928
7.728 | -34.2727 | -13.5027
7.7341 | -34.273 | -13.5126
7.7402 | -34.2734 | -13.5223
7.7463 | -34.2737 | -13.5321
7.7524 | -34.2739 | -13.5421
7.7585 | -34.2742 | -13.552
7.7646 | -34.2745 | -13.5615
7.7707 | -34.2748 | -13.5711
7.7768 | -34.275 | -13.5808
7.7829 | -34.2753 | -13.5907
7.789 | -34.2756 | -13.6002
7.7951 | -34.2759 | -13.6104
7.8012 | -34.2762 | -13.6205
7.8073 | -34.2765 | -13.6302
7.8135 | -34.2768 | -13.6401
7.8196 | -34.2771 | -13.6498
7.8257 | -34.2775 | -13.66
7.8318 | -34.2777 | -13.6702
7.8379 | -34.2781 | -13.6795
7.844 | -34.2783 | -13.6889
7.8501 | -34.2785 | -13.699
7.8562 | -34.2788 | -13.7094
7.8623 | -34.2792 | -13.7192
7.8684 | -34.2794 | -13.7287
7.8745 | -34.2797 | -13.7384
7.8806 | -34.28 | -13.7477
7.8867 | -34.2803 | -13.7573
7.8928 | -34.2807 | -13.7669
7.8989 | -34.281 | -13.7763
7.905 | -34.2813 | -13.7864
7.9111 | -34.2816 | -13.7963
7.9172 | -34.2819 | -13.8058
7.9233 | -34.2822 | -13.8157
7.9294 | -34.2825 | -13.8256
7.9355 | -34.2828 | -13.835
7.9416 | -34.283 | -13.8447
7.9477 | -34.2834 | -13.8541
7.9539 | -34.2837 | -13.8635
7.96 | -34.284 | -13.8735
7.9661 | -34.2843 | -13.8833
7.9722 | -34.2846 | -13.8931
7.9783 | -34.2849 | -13.9029
7.9844 | -34.2853 | -13.9131
7.9905 | -34.2855 | -13.9228
7.9966 | -34.2858 | -13.9326
8.0027 | -34.2861 | -13.9427
8.0088 | -34.2864 | -13.952
8.0149 | -34.2867 | -13.9615
8.021 | -34.287 | -13.9714
8.0271 | -34.2873 | -13.9816
8.0332 | -34.2875 | -13.9913
8.0393 | -34.2879 | -14.0013
8.0454 | -34.2882 | -14.0113
8.0515 | -34.2885 | -14.021
8.0576 | -34.2887 | -14.031
8.0637 | -34.289 | -14.0403
8.0698 | -34.2893 | -14.0497
8.0759 | -34.2896 | -14.0592
8.082 | -34.2899 | -14.0692
8.0881 | -34.2902 | -14.0788
8.0942 | -34.2905 | -14.0891
8.1004 | -34.2908 | -14.0993
8.1065 | -34.2911 | -14.1089
8.1126 | -34.2914 | -14.1186
8.1187 | -34.2917 | -14.1285
8.1248 | -34.292 | -14.1385
8.1309 | -34.2923 | -14.1483
8.137 | -34.2925 | -14.158
8.1431 | -34.2929 | -14.1679
8.1492 | -34.2932 | -14.1777
8.1553 | -34.2935 | -14.1875
8.1614 | -34.2937 | -14.1976
8.1675 | -34.294 | -14.2072
8.1736 | -34.2943 | -14.2172
8.1797 | -34.2946 | -14.2272
8.1858 | -34.2948 | -14.2369
8.1919 | -34.2951 | -14.2467
8.198 | -34.2954 | -14.2561
8.2041 | -34.2958 | -14.2662
8.2102 | -34.2961 | -14.2759
8.2163 | -34.2964 | -14.2857
8.2224 | -34.2967 | -14.2957
8.2285 | -34.297 | -14.3053
8.2346 | -34.2973 | -14.3149
8.2408 | -34.2976 | -14.3248
8.2469 | -34.2979 | -14.3345
8.253 | -34.2981 | -14.3442
8.2591 | -34.2984 | -14.3537
8.2652 | -34.2988 | -14.3633
8.2713 | -34.2992 | -14.3732
8.2774 | -34.2994 | -14.383
8.2835 | -34.2998 | -14.3932
8.2896 | -34.3 | -14.4025
8.2957 | -34.3003 | -14.4123
8.3018 | -34.3006 | -14.422
8.3079 | -34.301 | -14.4317
8.314 | -34.3013 | -14.4413
8.3201 | -34.3016 | -14.451
8.3262 | -34.3019 | -14.4607
8.3323 | -34.3022 | -14.4704
8.3384 | -34.3025 | -14.4807
8.3445 | -34.3028 | -14.49
8.3506 | -34.3031 | -14.5001
8.3567 | -34.3034 | -14.5101
8.3628 | -34.3037 | -14.5198
8.3689 | -34.304 | -14.5296
8.375 | -34.3042 | -14.5394
8.3812 | -34.3045 | -14.5492
8.3873 | -34.3048 | -14.5591
8.3934 | -34.3052 | -14.5689
8.3995 | -34.3054 | -14.5788
8.4056 | -34.3057 | -14.5887
8.4117 | -34.306 | -14.5987
8.4178 | -34.3064 | -14.608
8.4239 | -34.3067 | -14.6176
8.43 | -34.307 | -14.6273
8.4361 | -34.3073 | -14.6375
8.4422 | -34.3076 | -14.6472
8.4483 | -34.3078 | -14.6569
8.4544 | -34.3082 | -14.667
8.4605 | -34.3084 | -14.6766
8.4666 | -34.3087 | -14.6866
8.4727 | -34.309 | -14.6962
8.4788 | -34.3093 | -14.7059
8.4849 | -34.3096 | -14.716
8.491 | -34.3099 | -14.7255
8.4971 | -34.3102 | -14.7354
8.5032 | -34.3105 | -14.7452
8.5093 | -34.3108 | -14.755
8.5154 | -34.3111 | -14.7646
8.5215 | -34.3114 | -14.7743
8.5277 | -34.3117 | -14.7842
8.5338 | -34.312 | -14.7936
8.5399 | -34.3123 | -14.8037
8.546 | -34.3126 | -14.8132
8.5521 | -34.3129 | -14.8231
8.5582 | -34.3132 | -14.833
8.5643 | -34.3135 | -14.8425
8.5704 | -34.3138 | -14.852
8.5765 | -34.3141 | -14.8617
8.5826 | -34.3145 | -14.8713
8.5887 | -34.3148 | -14.8809
8.5948 | -34.3151 | -14.8903
8.6009 | -34.3154 | -14.8999
8.607 | -34.3157 | -14.9099
8.6131 | -34.3159 | -14.9199
8.6192 | -34.3162 | -14.9292
8.6253 | -34.3166 | -14.9391
8.6314 | -34.3169 | -14.9487
8.6375 | -34.3172 | -14.9587
8.6436 | -34.3176 | -14.9685
8.6497 | -34.3179 | -14.9784
8.6558 | -34.3181 | -14.988
8.6619 | -34.3184 | -14.9979
8.6681 | -34.3187 | -15.0076
8.6742 | -34.319 | -15.0176
8.6803 | -34.3194 | -15.0276
8.6864 | -34.3197 | -15.0376
8.6925 | -34.32 | -15.0473
8.6986 | -34.3203 | -15.0575
8.7047 | -34.3207 | -15.0673
8.7108 | -34.321 | -15.0772
8.7169 | -34.3213 | -15.0868
8.723 | -34.3216 | -15.0966
8.7291 | -34.3219 | -15.1062
8.7352 | -34.3223 | -15.1162
8.7413 | -34.3226 | -15.1259
8.7474 | -34.3229 | -15.1353
8.7535 | -34.3232 | -15.1452
8.7596 | -34.3236 | -15.1555
8.7657 | -34.3239 | -15.1652
8.7718 | -34.3242 | -15.1747
8.7779 | -34.3245 | -15.185
8.784 | -34.3247 | -15.1946
8.7901 | -34.325 | -15.2048
8.7962 | -34.3254 | -15.2147
8.8023 | -34.3258 | -15.2241
8.8084 | -34.3261 | -15.2341
8.8146 | -34.3264 | -15.2437
8.8207 | -34.3267 | -15.2535
8.8268 | -34.3271 | -15.2632
8.8329 | -34.3274 | -15.273
8.839 | -34.3278 | -15.2829
8.8451 | -34.3281 | -15.2924
8.8512 | -34.3284 | -15.3022
8.8573 | -34.3287 | -15.3118
8.8634 | -34.3291 | -15.3214
8.8695 | -34.3294 | -15.3313
8.8756 | -34.3298 | -15.3409
8.8817 | -34.3301 | -15.3509
8.8878 | -34.3303 | -15.3609
8.8939 | -34.3306 | -15.3703
8.9 | -34.331 | -15.3797
8.9061 | -34.3314 | -15.3895
8.9122 | -34.3317 | -15.3991
8.9183 | -34.332 | -15.4088
8.9244 | -34.3323 | -15.4184
8.9305 | -34.3327 | -15.428
8.9366 | -34.333 | -15.4376
8.9427 | -34.3334 | -15.4474
8.9488 | -34.3337 | -15.4571
8.955 | -34.3341 | -15.4668
8.9611 | -34.3344 | -15.4767
8.9672 | -34.3348 | -15.4868
8.9733 | -34.3351 | -15.4964
8.9794 | -34.3354 | -15.5062
8.9855 | -34.3358 | -15.5158
8.9916 | -34.3361 | -15.5254
8.9977 | -34.3364 | -15.535
9.0038 | -34.3368 | -15.5448
9.0099 | -34.3372 | -15.5546
9.016 | -34.3376 | -15.5644
9.0221 | -34.3379 | -15.5744
9.0282 | -34.3383 | -15.5843
9.0343 | -34.3386 | -15.5942
9.0404 | -34.3391 | -15.6038
9.0465 | -34.3394 | -15.6139
9.0526 | -34.3398 | -15.6234
9.0587 | -34.3402 | -15.6331
9.0648 | -34.3406 | -15.6427
9.0709 | -34.3409 | -15.6522
9.077 | -34.3412 | -15.6625
9.0831 | -34.3416 | -15.6723
9.0892 | -34.3419 | -15.682
9.0953 | -34.3422 | -15.6918
9.1015 | -34.3426 | -15.7014
9.1076 | -34.343 | -15.7112
9.1137 | -34.3434 | -15.7209
9.1198 | -34.3438 | -15.7303
9.1259 | -34.3441 | -15.7401
9.132 | -34.3445 | -15.7501
9.1381 | -34.3449 | -15.7598
9.1442 | -34.3454 | -15.7691
9.1503 | -34.3457 | -15.779
9.1564 | -34.346 | -15.7887
9.1625 | -34.3464 | -15.7984
9.1686 | -34.3468 | -15.8082
9.1747 | -34.3472 | -15.818
9.1808 | -34.3476 | -15.828
9.1869 | -34.3479 | -15.8379
9.193 | -34.3483 | -15.8478
9.1991 | -34.3486 | -15.8573
9.2052 | -34.349 | -15.8669
9.2113 | -34.3494 | -15.8768
9.2174 | -34.3498 | -15.8871
9.2235 | -34.3502 | -15.8966
9.2296 | -34.3506 | -15.9064
9.2357 | -34.3509 | -15.9162
9.2419 | -34.3513 | -15.9259
9.248 | -34.3517 | -15.9359
9.2541 | -34.3521 | -15.9458
9.2602 | -34.3524 | -15.9556
9.2663 | -34.3528 | -15.9652
9.2724 | -34.3532 | -15.975
9.2785 | -34.3536 | -15.9847
9.2846 | -34.354 | -15.9944
9.2907 | -34.3543 | -16.0041
9.2968 | -34.3547 | -16.014
9.3029 | -34.3551 | -16.0238
9.309 | -34.3555 | -16.0334
9.3151 | -34.3558 | -16.0431
9.3212 | -34.3563 | -16.053
9.3273 | -34.3567 | -16.0626
9.3334 | -34.3572 | -16.0721
9.3395 | -34.3576 | -16.0816
9.3456 | -34.358 | -16.0911
9.3517 | -34.3584 | -16.1005
9.3578 | -34.3588 | -16.1101
9.3639 | -34.3592 | -16.12
9.37 | -34.3596 | -16.1297
9.3761 | -34.36 | -16.1393
9.3822 | -34.3604 | -16.149
9.3884 | -34.3608 | -16.1588
9.3945 | -34.3612 | -16.1685
9.4006 | -34.3616 | -16.1783
9.4067 | -34.3619 | -16.1882
9.4128 | -34.3624 | -16.1978
9.4189 | -34.3628 | -16.2076
9.425 | -34.3632 | -16.2171
9.4311 | -34.3636 | -16.2269
9.4372 | -34.364 | -16.2366
9.4433 | -34.3644 | -16.2464
9.4494 | -34.3648 | -16.2561
9.4555 | -34.3652 | -16.2658
9.4616 | -34.3656 | -16.2755
9.4677 | -34.3661 | -16.2851
9.4738 | -34.3664 | -16.2951
9.4799 | -34.3668 | -16.3047
9.486 | -34.3671 | -16.3147
9.4921 | -34.3676 | -16.3249
9.4982 | -34.368 | -16.3342
9.5043 | -34.3684 | -16.3437
9.5104 | -34.3688 | -16.3533
9.5165 | -34.3692 | -16.3629
9.5226 | -34.3696 | -16.3728
9.5288 | -34.3701 | -16.3826
9.5349 | -34.3705 | -16.3922
9.541 | -34.3709 | -16.4019
9.5471 | -34.3713 | -16.4117
9.5532 | -34.3717 | -16.4214
9.5593 | -34.3722 | -16.4312
9.5654 | -34.3726 | -16.441
9.5715 | -34.373 | -16.451
9.5776 | -34.3734 | -16.461
9.5837 | -34.3738 | -16.4705
9.5898 | -34.3741 | -16.4803
9.5959 | -34.3745 | -16.4899
9.602 | -34.3749 | -16.4996
9.6081 | -34.3753 | -16.5095
9.6142 | -34.3758 | -16.519
9.6203 | -34.3761 | -16.5288
9.6264 | -34.3765 | -16.5381
9.6325 | -34.3769 | -16.548
9.6386 | -34.3773 | -16.5576
9.6447 | -34.3777 | -16.5672
9.6508 | -34.3781 | -16.577
9.6569 | -34.3785 | -16.5866
9.663 | -34.379 | -16.5964
9.6691 | -34.3794 | -16.6062
9.6753 | -34.3799 | -16.6159
9.6814 | -34.3803 | -16.6253
9.6875 | -34.3807 | -16.6348
9.6936 | -34.3811 | -16.6443
9.6997 | -34.3815 | -16.6542
9.7058 | -34.382 | -16.6639
9.7119 | -34.3824 | -16.6735
9.718 | -34.3828 | -16.6835
9.7241 | -34.3832 | -16.6933
9.7302 | -34.3835 | -16.703
9.7363 | -34.384 | -16.7125
9.7424 | -34.3844 | -16.722
9.7485 | -34.3848 | -16.7316
9.7546 | -34.3853 | -16.741
9.7607 | -34.3856 | -16.751
9.7668 | -34.3861 | -16.7611
9.7729 | -34.3865 | -16.7705
9.779 | -34.387 | -16.7802
9.7851 | -34.3873 | -16.7901
9.7912 | -34.3878 | -16.8003
9.7973 | -34.3881 | -16.8103
9.8034 | -34.3886 | -16.8201
9.8095 | -34.389 | -16.8298
9.8157 | -34.3894 | -16.8396
9.8218 | -34.3898 | -16.8489
9.8279 | -34.3903 | -16.8587
9.834 | -34.3907 | -16.8682
9.8401 | -34.3911 | -16.8778
9.8462 | -34.3916 | -16.8873
9.8523 | -34.392 | -16.8969
9.8584 | -34.3924 | -16.9069
9.8645 | -34.3928 | -16.9166
9.8706 | -34.3932 | -16.9261
9.8767 | -34.3936 | -16.9354
9.8828 | -34.3941 | -16.9447
9.8889 | -34.3945 | -16.9545
9.895 | -34.3949 | -16.9639
9.9011 | -34.3953 | -16.973
9.9072 | -34.3958 | -16.9824
9.9133 | -34.3962 | -16.9919
9.9194 | -34.3966 | -17.0013
9.9255 | -34.397 | -17.0111
9.9316 | -34.3975 | -17.0205
9.9377 | -34.3979 | -17.03
9.9438 | -34.3983 | -17.0395
9.9499 | -34.3987 | -17.0493
9.956 | -34.3992 | -17.0588
9.9622 | -34.3996 | -17.0683
9.9683 | -34.4 | -17.0779
9.9744 | -34.4005 | -17.0875
9.9805 | -34.4009 | -17.0971
9.9866 | -34.4013 | -17.1064
9.9927 | -34.4018 | -17.1163
9.9988 | -34.4022 | -17.1258
10.0049 | -34.4026 | -17.1354
10.011 | -34.4031 | -17.1449
10.0171 | -34.4035 | -17.1544
10.0232 | -34.4039 | -17.1637
10.0293 | -34.4044 | -17.173
10.0354 | -34.4049 | -17.1823
10.0415 | -34.4053 | -17.1918
10.0476 | -34.4057 | -17.2018
10.0537 | -34.4061 | -17.2116
10.0598 | -34.4066 | -17.2213
10.0659 | -34.4069 | -17.2307
10.072 | -34.4074 | -17.2402
10.0781 | -34.4079 | -17.2498
10.0842 | -34.4083 | -17.2595
10.0903 | -34.4087 | -17.2694
10.0964 | -34.4092 | -17.2785
10.1026 | -34.4097 | -17.2881
10.1087 | -34.4101 | -17.2978
10.1148 | -34.4105 | -17.3075
10.1209 | -34.411 | -17.3169
10.127 | -34.4114 | -17.3262
10.1331 | -34.4119 | -17.3356
10.1392 | -34.4123 | -17.345
10.1453 | -34.4128 | -17.3548
10.1514 | -34.4132 | -17.3641
10.1575 | -34.4136 | -17.3735
10.1636 | -34.4141 | -17.3823
10.1697 | -34.4145 | -17.3919
10.1758 | -34.4149 | -17.4012
10.1819 | -34.4154 | -17.4107
10.188 | -34.4158 | -17.42
10.1941 | -34.4162 | -17.4296
10.2002 | -34.4165 | -17.4384
10.2063 | -34.417 | -17.4481
10.2124 | -34.4174 | -17.4578
10.2185 | -34.4179 | -17.4673
10.2246 | -34.4183 | -17.4767
10.2307 | -34.4187 | -17.4864
10.2368 | -34.4192 | -17.4949
10.2429 | -34.4196 | -17.5042
10.2491 | -34.42 | -17.5136
10.2552 | -34.4205 | -17.5233
10.2613 | -34.421 | -17.5331
10.2674 | -34.4212 | -17.5425
10.2735 | -34.4216 | -17.5523
10.2796 | -34.4221 | -17.562
10.2857 | -34.4225 | -17.5716
10.2918 | -34.4229 | -17.5812
10.2979 | -34.4234 | -17.5907
10.304 | -34.4239 | -17.6003
10.3101 | -34.4243 | -17.6099
10.3162 | -34.4248 | -17.6193
10.3223 | -34.4252 | -17.6288
10.3284 | -34.4256 | -17.6382
10.3345 | -34.4261 | -17.6475
10.3406 | -34.4265 | -17.6567
10.3467 | -34.4269 | -17.6661
10.3528 | -34.4273 | -17.6752
10.3589 | -34.4277 | -17.6848
10.365 | -34.4281 | -17.6946
10.3711 | -34.4285 | -17.7043
10.3772 | -34.4289 | -17.7137
10.3833 | -34.4293 | -17.7231
10.3895 | -34.4297 | -17.7323
10.3956 | -34.4302 | -17.7415
10.4017 | -34.4306 | -17.751
10.4078 | -34.431 | -17.7604
10.4139 | -34.4315 | -17.77
10.42 | -34.4319 | -17.7795
10.4261 | -34.4324 | -17.7888
10.4322 | -34.4328 | -17.798
10.4383 | -34.4332 | -17.8076
10.4444 | -34.4336 | -17.8169
10.4505 | -34.4341 | -17.8264
10.4566 | -34.4344 | -17.8358
10.4627 | -34.4349 | -17.8453
10.4688 | -34.4353 | -17.8543
10.4749 | -34.4357 | -17.8637
10.481 | -34.4361 | -17.8733
10.4871 | -34.4365 | -17.8828
10.4932 | -34.437 | -17.8921
10.4993 | -34.4375 | -17.9016
10.5054 | -34.4379 | -17.9109
10.5115 | -34.4383 | -17.9203
10.5176 | -34.4388 | -17.9293
10.5237 | -34.4393 | -17.9386
10.5298 | -34.4396 | -17.9477
10.536 | -34.4401 | -17.9573
10.5421 | -34.4405 | -17.9662
10.5482 | -34.4409 | -17.9758
10.5543 | -34.4414 | -17.9852
10.5604 | -34.4418 | -17.9942
10.5665 | -34.4422 | -18.0033
10.5726 | -34.4426 | -18.0125
10.5787 | -34.4431 | -18.0219
10.5848 | -34.4436 | -18.0308
10.5909 | -34.444 | -18.0403
10.597 | -34.4443 | -18.0497
10.6031 | -34.4448 | -18.0593
10.6092 | -34.4452 | -18.0689
10.6153 | -34.4456 | -18.0785
10.6214 | -34.4461 | -18.0876
10.6275 | -34.4465 | -18.0967
10.6336 | -34.4469 | -18.1065
10.6397 | -34.4473 | -18.1158
10.6458 | -34.4477 | -18.1251
10.6519 | -34.4481 | -18.1345
10.658 | -34.4485 | -18.1436
10.6641 | -34.4489 | -18.1529
10.6702 | -34.4494 | -18.1626
10.6764 | -34.4499 | -18.172
10.6825 | -34.4503 | -18.1814
10.6886 | -34.4508 | -18.1909
10.6947 | -34.4512 | -18.2001
10.7008 | -34.4516 | -18.2093
10.7069 | -34.452 | -18.2188
10.713 | -34.4525 | -18.228
10.7191 | -34.4529 | -18.2376
10.7252 | -34.4534 | -18.2469
10.7313 | -34.4538 | -18.256
10.7374 | -34.4542 | -18.2651
10.7435 | -34.4546 | -18.2747
10.7496 | -34.455 | -18.2838
10.7557 | -34.4554 | -18.293
10.7618 | -34.4559 | -18.3023
10.7679 | -34.4563 | -18.3115
10.774 | -34.4567 | -18.3208
10.7801 | -34.4571 | -18.3304
10.7862 | -34.4575 | -18.3395
10.7923 | -34.458 | -18.3491
10.7984 | -34.4584 | -18.3583
10.8045 | -34.4588 | -18.3677
10.8106 | -34.4592 | -18.3771
10.8168 | -34.4597 | -18.3863
10.8229 | -34.4601 | -18.3957
10.829 | -34.4605 | -18.4051
10.8351 | -34.4609 | -18.4143
10.8412 | -34.4613 | -18.4235
10.8473 | -34.4617 | -18.4325
10.8534 | -34.4621 | -18.442
10.8595 | -34.4625 | -18.451
10.8656 | -34.463 | -18.4603
10.8717 | -34.4634 | -18.4693
10.8778 | -34.4638 | -18.4785
10.8839 | -34.4642 | -18.4877
10.89 | -34.4646 | -18.4968
10.8961 | -34.4651 | -18.5061
10.9022 | -34.4655 | -18.5151
10.9083 | -34.4659 | -18.5247
10.9144 | -34.4664 | -18.5339
10.9205 | -34.4668 | -18.5432
10.9266 | -34.4673 | -18.5528
10.9327 | -34.4678 | -18.5622
10.9388 | -34.4681 | -18.5714
10.9449 | -34.4686 | -18.5807
10.951 | -34.469 | -18.59
10.9571 | -34.4694 | -18.5993
10.9633 | -34.4698 | -18.609
10.9694 | -34.4702 | -18.6185
10.9755 | -34.4707 | -18.6278
10.9816 | -34.4711 | -18.6371
10.9877 | -34.4715 | -18.6464
10.9938 | -34.4719 | -18.6559
10.9999 | -34.4724 | -18.6651
11.006 | -34.4728 | -18.6743
11.0121 | -34.4732 | -18.6839
11.0182 | -34.4737 | -18.6928
11.0243 | -34.4741 | -18.7022
11.0304 | -34.4745 | -18.7114
11.0365 | -34.4749 | -18.7205
11.0426 | -34.4753 | -18.7295
11.0487 | -34.4758 | -18.7388
11.0548 | -34.4762 | -18.748
11.0609 | -34.4767 | -18.7574
11.067 | -34.477 | -18.7668
11.0731 | -34.4775 | -18.7761
11.0792 | -34.4779 | -18.7852
11.0853 | -34.4783 | -18.7941
11.0914 | -34.4787 | -18.8033
11.0975 | -34.4792 | -18.8128
11.1037 | -34.4796 | -18.8218
11.1098 | -34.48 | -18.8313
11.1159 | -34.4805 | -18.8406
11.122 | -34.4809 | -18.8502
11.1281 | -34.4813 | -18.8592
11.1342 | -34.4817 | -18.868
11.1403 | -34.4821 | -18.8776
11.1464 | -34.4826 | -18.8867
11.1525 | -34.483 | -18.896
11.1586 | -34.4834 | -18.9057
11.1647 | -34.4839 | -18.9147
11.1708 | -34.4843 | -18.9239
11.1769 | -34.4847 | -18.9325
11.183 | -34.4851 | -18.9416
11.1891 | -34.4856 | -18.9511
11.1952 | -34.486 | -18.9607
11.2013 | -34.4864 | -18.9698
11.2074 | -34.4868 | -18.9787
11.2135 | -34.4871 | -18.988
11.2196 | -34.4875 | -18.997
11.2257 | -34.4879 | -19.0062
11.2318 | -34.4883 | -19.0155
11.2379 | -34.4887 | -19.0246
11.244 | -34.4891 | -19.0338
11.2502 | -34.4895 | -19.043
11.2563 | -34.49 | -19.0522
11.2624 | -34.4903 | -19.0615
11.2685 | -34.4907 | -19.0704
11.2746 | -34.4911 | -19.0793
11.2807 | -34.4915 | -19.0888
11.2868 | -34.4919 | -19.0983
11.2929 | -34.4924 | -19.107
11.299 | -34.4928 | -19.1164
11.3051 | -34.4932 | -19.1255
11.3112 | -34.4936 | -19.1349
11.3173 | -34.4941 | -19.1441
11.3234 | -34.4944 | -19.1534
11.3295 | -34.4949 | -19.1624
11.3356 | -34.4953 | -19.1716
11.3417 | -34.4957 | -19.1808
11.3478 | -34.4961 | -19.1899
11.3539 | -34.4964 | -19.199
11.36 | -34.4969 | -19.2081
11.3661 | -34.4973 | -19.2172
11.3722 | -34.4977 | -19.2261
11.3783 | -34.4981 | -19.2353
11.3844 | -34.4985 | -19.2445
11.3906 | -34.4989 | -19.2539
11.3967 | -34.4993 | -19.2629
11.4028 | -34.4997 | -19.2722
11.4089 | -34.5001 | -19.2811
11.415 | -34.5005 | -19.2902
11.4211 | -34.5009 | -19.2991
11.4272 | -34.5012 | -19.3083
11.4333 | -34.5016 | -19.3175
11.4394 | -34.5021 | -19.3269
11.4455 | -34.5024 | -19.3361
11.4516 | -34.5028 | -19.3451
11.4577 | -34.5033 | -19.3546
11.4638 | -34.5038 | -19.3633
11.4699 | -34.5041 | -19.3728
11.476 | -34.5045 | -19.3824
11.4821 | -34.505 | -19.3914
11.4882 | -34.5054 | -19.4009
11.4943 | -34.5058 | -19.4099
11.5004 | -34.5062 | -19.4191
11.5065 | -34.5067 | -19.4282
11.5126 | -34.5071 | -19.4371
11.5187 | -34.5076 | -19.4462
11.5248 | -34.508 | -19.4554
11.5309 | -34.5084 | -19.4645
11.5371 | -34.5088 | -19.4737
11.5432 | -34.5091 | -19.4831
11.5493 | -34.5095 | -19.4924
11.5554 | -34.5099 | -19.5013
11.5615 | -34.5104 | -19.5106
11.5676 | -34.5108 | -19.5199
11.5737 | -34.5112 | -19.5288
11.5798 | -34.5116 | -19.5377
11.5859 | -34.5121 | -19.5467
11.592 | -34.5125 | -19.5559
11.5981 | -34.5129 | -19.5649
11.6042 | -34.5133 | -19.5741
11.6103 | -34.5137 | -19.583
11.6164 | -34.5142 | -19.5919
11.6225 | -34.5145 | -19.6009
11.6286 | -34.515 | -19.6101
11.6347 | -34.5154 | -19.6196
11.6408 | -34.5158 | -19.6286
11.6469 | -34.5163 | -19.6379
11.653 | -34.5167 | -19.6466
11.6591 | -34.5171 | -19.6558
11.6652 | -34.5175 | -19.665
11.6713 | -34.5179 | -19.674
11.6775 | -34.5184 | -19.683
11.6836 | -34.5188 | -19.6922
11.6897 | -34.5193 | -19.7011
11.6958 | -34.5197 | -19.7105
11.7019 | -34.5201 | -19.7197
11.708 | -34.5205 | -19.729
11.7141 | -34.521 | -19.738
11.7202 | -34.5214 | -19.7468
11.7263 | -34.5218 | -19.756
11.7324 | -34.5222 | -19.7652
11.7385 | -34.5226 | -19.7744
11.7446 | -34.523 | -19.7836
11.7507 | -34.5234 | -19.7926
11.7568 | -34.5238 | -19.8018
11.7629 | -34.5243 | -19.8108
11.769 | -34.5247 | -19.82
11.7751 | -34.5251 | -19.8288
11.7812 | -34.5256 | -19.8376
11.7873 | -34.526 | -19.847
11.7934 | -34.5264 | -19.8561
11.7995 | -34.5268 | -19.8648
11.8056 | -34.5272 | -19.8742
11.8117 | -34.5276 | -19.8833
11.8178 | -34.528 | -19.8923
11.824 | -34.5285 | -19.9011
11.8301 | -34.529 | -19.9097
11.8362 | -34.5294 | -19.9185
11.8423 | -34.5297 | -19.9275
11.8484 | -34.5302 | -19.9364
11.8545 | -34.5306 | -19.9456
11.8606 | -34.531 | -19.9547
11.8667 | -34.5315 | -19.9638
11.8728 | -34.5319 | -19.9727
11.8789 | -34.5323 | -19.9821
11.885 | -34.5326 | -19.9911
11.8911 | -34.5331 | -20.0002
11.8972 | -34.5336 | -20.0089
11.9033 | -34.534 | -20.0176
11.9094 | -34.5343 | -20.0262
11.9155 | -34.5347 | -20.0352
11.9216 | -34.5352 | -20.044
11.9277 | -34.5357 | -20.0527
11.9338 | -34.536 | -20.0618
11.9399 | -34.5365 | -20.071
11.946 | -34.5369 | -20.08
11.9521 | -34.5373 | -20.0889
11.9582 | -34.5378 | -20.0977
11.9644 | -34.5382 | -20.1066
11.9705 | -34.5387 | -20.1158
11.9766 | -34.5391 | -20.1248
11.9827 | -34.5395 | -20.1337
11.9888 | -34.5399 | -20.1424
11.9949 | -34.5403 | -20.1512
12.001 | -34.5408 | -20.1602
12.0071 | -34.5412 | -20.169
12.0132 | -34.5416 | -20.1777
12.0193 | -34.542 | -20.1869
12.0254 | -34.5425 | -20.1956
12.0315 | -34.5429 | -20.2044
12.0376 | -34.5433 | -20.2131
12.0437 | -34.5437 | -20.2219
12.0498 | -34.5441 | -20.231
12.0559 | -34.5446 | -20.24
12.062 | -34.545 | -20.2488
12.0681 | -34.5454 | -20.2577
12.0742 | -34.5459 | -20.2668
12.0803 | -34.5463 | -20.2755
12.0864 | -34.5467 | -20.2843
12.0925 | -34.5472 | -20.2933
12.0986 | -34.5476 | -20.3019
12.1047 | -34.5481 | -20.3108
12.1109 | -34.5485 | -20.3197
12.117 | -34.549 | -20.3285
12.1231 | -34.5494 | -20.3375
12.1292 | -34.5498 | -20.3466
12.1353 | -34.5503 | -20.3552
12.1414 | -34.5507 | -20.3642
12.1475 | -34.5512 | -20.3731
12.1536 | -34.5516 | -20.3821
12.1597 | -34.552 | -20.391
12.1658 | -34.5524 | -20.3999
12.1719 | -34.553 | -20.4089
12.178 | -34.5534 | -20.4179
12.1841 | -34.5538 | -20.427
12.1902 | -34.5543 | -20.4359
12.1963 | -34.5547 | -20.4448
12.2024 | -34.5552 | -20.4534
12.2085 | -34.5556 | -20.4624
12.2146 | -34.556 | -20.4712
12.2207 | -34.5564 | -20.4801
12.2268 | -34.5568 | -20.4888
12.2329 | -34.5573 | -20.4977
12.239 | -34.5577 | -20.5062
12.2451 | -34.5582 | -20.5148
12.2513 | -34.5586 | -20.5238
12.2574 | -34.5591 | -20.5329
12.2635 | -34.5595 | -20.5413
12.2696 | -34.5599 | -20.55
12.2757 | -34.5604 | -20.5586
12.2818 | -34.5608 | -20.5672
12.2879 | -34.5612 | -20.5758
12.294 | -34.5617 | -20.5844
12.3001 | -34.5621 | -20.5933
12.3062 | -34.5626 | -20.6021
12.3123 | -34.5631 | -20.6109
12.3184 | -34.5635 | -20.6194
12.3245 | -34.564 | -20.6283
12.3306 | -34.5644 | -20.6369
12.3367 | -34.5648 | -20.6454
12.3428 | -34.5653 | -20.6539
12.3489 | -34.5658 | -20.6621
12.355 | -34.5662 | -20.6709
12.3611 | -34.5666 | -20.6797
12.3672 | -34.5671 | -20.6883
12.3733 | -34.5675 | -20.6973
12.3794 | -34.5679 | -20.7055
12.3855 | -34.5683 | -20.7141
12.3916 | -34.5688 | -20.7229
12.3978 | -34.5692 | -20.7319
12.4039 | -34.5696 | -20.7406
12.41 | -34.5701 | -20.7495
12.4161 | -34.5706 | -20.7579
12.4222 | -34.571 | -20.7665
12.4283 | -34.5715 | -20.7753
12.4344 | -34.5719 | -20.784
12.4405 | -34.5723 | -20.7928
12.4466 | -34.5728 | -20.8014
12.4527 | -34.5732 | -20.8101
12.4588 | -34.5737 | -20.8188
12.4649 | -34.5741 | -20.8277
12.471 | -34.5746 | -20.8368
12.4771 | -34.575 | -20.8449
12.4832 | -34.5755 | -20.8535
12.4893 | -34.5759 | -20.8619
12.4954 | -34.5764 | -20.8704
12.5015 | -34.5769 | -20.8789
12.5076 | -34.5773 | -20.8878
12.5137 | -34.5777 | -20.8969
12.5198 | -34.5781 | -20.9052
12.5259 | -34.5786 | -20.9137
12.532 | -34.5791 | -20.9226
12.5382 | -34.5795 | -20.9311
12.5443 | -34.5799 | -20.9399
12.5504 | -34.5804 | -20.9486
12.5565 | -34.5809 | -20.9571
12.5626 | -34.5813 | -20.9661
12.5687 | -34.5818 | -20.9748
12.5748 | -34.5822 | -20.9834
12.5809 | -34.5826 | -20.9916
12.587 | -34.5831 | -21.0006
12.5931 | -34.5836 | -21.0092
12.5992 | -34.584 | -21.0182
12.6053 | -34.5844 | -21.0267
12.6114 | -34.5849 | -21.0353
12.6175 | -34.5854 | -21.0441
12.6236 | -34.5858 | -21.0524
12.6297 | -34.5863 | -21.0608
12.6358 | -34.5868 | -21.0692
12.6419 | -34.5873 | -21.0781
12.648 | -34.5877 | -21.0869
12.6541 | -34.5882 | -21.0956
12.6602 | -34.5886 | -21.1037
12.6663 | -34.589 | -21.1121
12.6724 | -34.5895 | -21.1208
12.6785 | -34.59 | -21.1295
12.6847 | -34.5904 | -21.138
12.6908 | -34.5909 | -21.1466
12.6969 | -34.5914 | -21.1551
12.703 | -34.5918 | -21.1638
12.7091 | -34.5923 | -21.1725
12.7152 | -34.5927 | -21.181
12.7213 | -34.5931 | -21.1897
12.7274 | -34.5936 | -21.1983
12.7335 | -34.594 | -21.2064
12.7396 | -34.5945 | -21.2146
12.7457 | -34.5949 | -21.2233
12.7518 | -34.5954 | -21.2318
12.7579 | -34.5958 | -21.2408
12.764 | -34.5962 | -21.2494
12.7701 | -34.5966 | -21.2575
12.7762 | -34.5971 | -21.2661
12.7823 | -34.5975 | -21.2747
12.7884 | -34.5979 | -21.2832
12.7945 | -34.5984 | -21.2915
12.8006 | -34.5988 | -21.2994
12.8067 | -34.5993 | -21.3082
12.8128 | -34.5997 | -21.3164
12.8189 | -34.6001 | -21.3246
12.8251 | -34.6005 | -21.3328
12.8312 | -34.6009 | -21.3413
12.8373 | -34.6013 | -21.3498
12.8434 | -34.6018 | -21.3584
12.8495 | -34.6022 | -21.3668
12.8556 | -34.6027 | -21.3752
12.8617 | -34.6031 | -21.3836
12.8678 | -34.6035 | -21.3922
12.8739 | -34.604 | -21.4005
12.88 | -34.6044 | -21.409
12.8861 | -34.6048 | -21.4175
12.8922 | -34.6053 | -21.426
12.8983 | -34.6058 | -21.4341
12.9044 | -34.6062 | -21.4425
12.9105 | -34.6067 | -21.4509
12.9166 | -34.6072 | -21.4597
12.9227 | -34.6076 | -21.4686
12.9288 | -34.6081 | -21.4773
12.9349 | -34.6086 | -21.4861
12.941 | -34.6091 | -21.4946
12.9471 | -34.6096 | -21.503
12.9532 | -34.6101 | -21.5116
12.9593 | -34.6105 | -21.5202
12.9654 | -34.6109 | -21.5288
12.9716 | -34.6113 | -21.5374
12.9777 | -34.6117 | -21.5455
12.9838 | -34.6122 | -21.5538
12.9899 | -34.6126 | -21.5624
12.996 | -34.613 | -21.5711
13.0021 | -34.6134 | -21.5796
13.0082 | -34.6139 | -21.588
13.0143 | -34.6143 | -21.596
13.0204 | -34.6147 | -21.6041
13.0265 | -34.6152 | -21.6122
13.0326 | -34.6156 | -21.6209
13.0387 | -34.6161 | -21.6296
13.0448 | -34.6165 | -21.638
13.0509 | -34.6169 | -21.6463
13.057 | -34.6174 | -21.6544
13.0631 | -34.618 | -21.6627
13.0692 | -34.6184 | -21.671
13.0753 | -34.6189 | -21.6795
13.0814 | -34.6193 | -21.6877
13.0875 | -34.6198 | -21.6957
13.0936 | -34.6202 | -21.704
13.0997 | -34.6207 | -21.7121
13.1058 | -34.6211 | -21.7199
13.112 | -34.6215 | -21.7283
13.1181 | -34.622 | -21.7367
13.1242 | -34.6224 | -21.7451
13.1303 | -34.6229 | -21.7539
13.1364 | -34.6233 | -21.7621
13.1425 | -34.6237 | -21.7703
13.1486 | -34.6242 | -21.7786
13.1547 | -34.6246 | -21.7865
13.1608 | -34.6251 | -21.7948
13.1669 | -34.6255 | -21.8031
13.173 | -34.626 | -21.8114
13.1791 | -34.6264 | -21.8198
13.1852 | -34.6268 | -21.8285
13.1913 | -34.6272 | -21.8368
13.1974 | -34.6277 | -21.8451
13.2035 | -34.6281 | -21.8534
13.2096 | -34.6286 | -21.8617
13.2157 | -34.629 | -21.8699
13.2218 | -34.6295 | -21.8784
13.2279 | -34.63 | -21.8868
13.234 | -34.6304 | -21.8952
13.2401 | -34.6308 | -21.9035
13.2462 | -34.6313 | -21.9117
13.2524 | -34.6317 | -21.92
13.2585 | -34.6322 | -21.9283
13.2646 | -34.6326 | -21.9369
13.2707 | -34.6331 | -21.9452
13.2768 | -34.6335 | -21.954
13.2829 | -34.6339 | -21.9624
13.289 | -34.6344 | -21.9707
13.2951 | -34.6348 | -21.979
13.3012 | -34.6353 | -21.9874
13.3073 | -34.6357 | -21.996
13.3134 | -34.6361 | -22.0042
13.3195 | -34.6365 | -22.0124
13.3256 | -34.637 | -22.0208
13.3317 | -34.6375 | -22.0292
13.3378 | -34.6379 | -22.0376
13.3439 | -34.6383 | -22.0459
13.35 | -34.6388 | -22.0543
13.3561 | -34.6392 | -22.0629
13.3622 | -34.6396 | -22.0709
13.3683 | -34.6401 | -22.0791
13.3744 | -34.6406 | -22.0873
13.3805 | -34.641 | -22.0958
13.3866 | -34.6414 | -22.104
13.3927 | -34.6418 | -22.1121
13.3989 | -34.6422 | -22.1202
13.405 | -34.6427 | -22.1284
13.4111 | -34.6432 | -22.1366
13.4172 | -34.6436 | -22.1447
13.4233 | -34.644 | -22.1532
13.4294 | -34.6444 | -22.1612
13.4355 | -34.6449 | -22.1699
13.4416 | -34.6453 | -22.1783
13.4477 | -34.6457 | -22.1867
13.4538 | -34.6461 | -22.1948
13.4599 | -34.6466 | -22.2032
13.466 | -34.647 | -22.2119
13.4721 | -34.6474 | -22.2199
13.4782 | -34.6479 | -22.2283
13.4843 | -34.6484 | -22.2366
13.4904 | -34.6488 | -22.2442
13.4965 | -34.6492 | -22.2529
13.5026 | -34.6496 | -22.2612
13.5087 | -34.65 | -22.2691
13.5148 | -34.6504 | -22.2775
13.5209 | -34.6509 | -22.2855
13.527 | -34.6513 | -22.2941
13.5331 | -34.6517 | -22.3023
13.5393 | -34.6522 | -22.3105
13.5454 | -34.6526 | -22.3189
13.5515 | -34.653 | -22.3271
13.5576 | -34.6535 | -22.3356
13.5637 | -34.6539 | -22.3439
13.5698 | -34.6544 | -22.3524
13.5759 | -34.6547 | -22.3603
13.582 | -34.6551 | -22.3687
13.5881 | -34.6555 | -22.377
13.5942 | -34.656 | -22.3853
13.6003 | -34.6564 | -22.3937
13.6064 | -34.6569 | -22.402
13.6125 | -34.6573 | -22.4105
13.6186 | -34.6578 | -22.4187
13.6247 | -34.6582 | -22.427
13.6308 | -34.6587 | -22.4351
13.6369 | -34.6591 | -22.4434
13.643 | -34.6596 | -22.4518
13.6491 | -34.66 | -22.4604
13.6552 | -34.6604 | -22.4685
13.6613 | -34.6609 | -22.4769
13.6674 | -34.6613 | -22.4853
13.6735 | -34.6617 | -22.4937
13.6796 | -34.6622 | -22.5018
13.6858 | -34.6625 | -22.5103
13.6919 | -34.663 | -22.5187
13.698 | -34.6634 | -22.527
13.7041 | -34.6638 | -22.5356
13.7102 | -34.6643 | -22.544
13.7163 | -34.6648 | -22.5526
13.7224 | -34.6652 | -22.5606
13.7285 | -34.6657 | -22.5692
13.7346 | -34.6661 | -22.5774
13.7407 | -34.6665 | -22.5858
13.7468 | -34.667 | -22.5941
13.7529 | -34.6673 | -22.6022
13.759 | -34.6678 | -22.6107
13.7651 | -34.6682 | -22.619
13.7712 | -34.6686 | -22.6276
13.7773 | -34.669 | -22.636
13.7834 | -34.6695 | -22.6443
13.7895 | -34.67 | -22.6526
13.7956 | -34.6704 | -22.661
13.8017 | -34.6709 | -22.6692
13.8078 | -34.6713 | -22.6775
13.8139 | -34.6717 | -22.6856
13.82 | -34.6721 | -22.6936
13.8262 | -34.6726 | -22.7023
13.8323 | -34.673 | -22.7108
13.8384 | -34.6734 | -22.7191
13.8445 | -34.6739 | -22.7276
13.8506 | -34.6743 | -22.7361
13.8567 | -34.6748 | -22.7446
13.8628 | -34.6752 | -22.753
13.8689 | -34.6756 | -22.7613
13.875 | -34.6761 | -22.7693
13.8811 | -34.6765 | -22.7775
13.8872 | -34.677 | -22.7858
13.8933 | -34.6774 | -22.7939
13.8994 | -34.6778 | -22.8021
13.9055 | -34.6783 | -22.8103
13.9116 | -34.6787 | -22.8184
13.9177 | -34.6792 | -22.8264
13.9238 | -34.6797 | -22.8347
13.9299 | -34.6801 | -22.8432
13.936 | -34.6805 | -22.8513
13.9421 | -34.6809 | -22.8596
13.9482 | -34.6814 | -22.8679
13.9543 | -34.6819 | -22.8761
13.9604 | -34.6823 | -22.8845
13.9665 | -34.6827 | -22.8929
13.9727 | -34.6831 | -22.901
13.9788 | -34.6835 | -22.9095
13.9849 | -34.6839 | -22.9178
13.991 | -34.6844 | -22.9264
13.9971 | -34.6848 | -22.9351
14.0032 | -34.6852 | -22.9435
14.0093 | -34.6856 | -22.9517
14.0154 | -34.686 | -22.9602
14.0215 | -34.6864 | -22.9682
14.0276 | -34.6869 | -22.9764
14.0337 | -34.6874 | -22.985
14.0398 | -34.6878 | -22.9929
14.0459 | -34.6882 | -23.0014
14.052 | -34.6886 | -23.0098
14.0581 | -34.689 | -23.0177
14.0642 | -34.6894 | -23.0262
14.0703 | -34.6899 | -23.0344
14.0764 | -34.6903 | -23.043
14.0825 | -34.6908 | -23.0513
14.0886 | -34.6912 | -23.0591
14.0947 | -34.6916 | -23.0672
14.1008 | -34.692 | -23.0759
14.1069 | -34.6925 | -23.0841
14.1131 | -34.6929 | -23.0927
14.1192 | -34.6933 | -23.1007
14.1253 | -34.6937 | -23.1087
14.1314 | -34.6942 | -23.1169
14.1375 | -34.6946 | -23.1255
14.1436 | -34.6951 | -23.1337
14.1497 | -34.6956 | -23.1418
14.1558 | -34.696 | -23.1502
14.1619 | -34.6964 | -23.1585
14.168 | -34.6969 | -23.1667
14.1741 | -34.6973 | -23.1749
14.1802 | -34.6977 | -23.1831
14.1863 | -34.6982 | -23.1913
14.1924 | -34.6986 | -23.1996
14.1985 | -34.699 | -23.2078
14.2046 | -34.6995 | -23.2158
14.2107 | -34.6999 | -23.2237
14.2168 | -34.7004 | -23.2324
14.2229 | -34.7008 | -23.2407
14.229 | -34.7013 | -23.249
14.2351 | -34.7017 | -23.257
14.2412 | -34.7021 | -23.2651
14.2473 | -34.7026 | -23.2734
14.2534 | -34.7031 | -23.2815
14.2596 | -34.7035 | -23.2898
14.2657 | -34.7039 | -23.2981
14.2718 | -34.7043 | -23.3065
14.2779 | -34.7048 | -23.3145
14.284 | -34.7052 | -23.3227
14.2901 | -34.7056 | -23.3308
14.2962 | -34.7061 | -23.3389
14.3023 | -34.7065 | -23.3473
14.3084 | -34.707 | -23.3558
14.3145 | -34.7074 | -23.3641
14.3206 | -34.7078 | -23.3722
14.3267 | -34.7082 | -23.3803
14.3328 | -34.7087 | -23.3885
14.3389 | -34.7091 | -23.3964
14.345 | -34.7095 | -23.4048
14.3511 | -34.7099 | -23.4129
14.3572 | -34.7103 | -23.4214
14.3633 | -34.7108 | -23.4295
14.3694 | -34.7112 | -23.4377
14.3755 | -34.7116 | -23.4455
14.3816 | -34.712 | -23.4537
14.3877 | -34.7125 | -23.4617
14.3938 | -34.7129 | -23.4701
14.4 | -34.7134 | -23.4781
14.4061 | -34.7138 | -23.4862
14.4122 | -34.7142 | -23.4946
14.4183 | -34.7146 | -23.5027
14.4244 | -34.715 | -23.5107
14.4305 | -34.7154 | -23.5192
14.4366 | -34.7159 | -23.5274
14.4427 | -34.7164 | -23.5359
14.4488 | -34.7167 | -23.5438
14.4549 | -34.7172 | -23.552
14.461 | -34.7177 | -23.5602
14.4671 | -34.7181 | -23.5686
14.4732 | -34.7186 | -23.5765
14.4793 | -34.719 | -23.5854
14.4854 | -34.7195 | -23.5932
14.4915 | -34.7199 | -23.6012
14.4976 | -34.7204 | -23.6095
14.5037 | -34.7209 | -23.6179
14.5098 | -34.7213 | -23.6259
14.5159 | -34.7218 | -23.6342
14.522 | -34.7222 | -23.6424
14.5281 | -34.7227 | -23.6503
14.5342 | -34.7231 | -23.6586
14.5403 | -34.7235 | -23.6668
14.5465 | -34.724 | -23.6753
14.5526 | -34.7245 | -23.6839
14.5587 | -34.7249 | -23.692
14.5648 | -34.7253 | -23.7
14.5709 | -34.7258 | -23.7083
14.577 | -34.7262 | -23.7166
14.5831 | -34.7267 | -23.7249
14.5892 | -34.7272 | -23.7328
14.5953 | -34.7276 | -23.7407
14.6014 | -34.7281 | -23.749
14.6075 | -34.7286 | -23.7569
14.6136 | -34.729 | -23.7648
14.6197 | -34.7294 | -23.773
14.6258 | -34.7298 | -23.7812
14.6319 | -34.7303 | -23.7891
14.638 | -34.7308 | -23.7973
14.6441 | -34.7312 | -23.805
14.6502 | -34.7316 | -23.813
14.6563 | -34.732 | -23.821
14.6624 | -34.7325 | -23.829
14.6685 | -34.7329 | -23.8368
14.6746 | -34.7334 | -23.8449
14.6807 | -34.7338 | -23.8531
14.6869 | -34.7343 | -23.8611
14.693 | -34.7347 | -23.8691
14.6991 | -34.7351 | -23.877
14.7052 | -34.7356 | -23.885
14.7113 | -34.7361 | -23.8934
14.7174 | -34.7365 | -23.9015
14.7235 | -34.737 | -23.9095
14.7296 | -34.7374 | -23.9175
14.7357 | -34.7379 | -23.9254
14.7418 | -34.7383 | -23.9332
14.7479 | -34.7388 | -23.9413
14.754 | -34.7392 | -23.9494
14.7601 | -34.7397 | -23.9576
14.7662 | -34.7402 | -23.9657
14.7723 | -34.7406 | -23.9737
14.7784 | -34.741 | -23.9819
14.7845 | -34.7415 | -23.9901
14.7906 | -34.7419 | -23.9983
14.7967 | -34.7424 | -24.0062
14.8028 | -34.7428 | -24.014
14.8089 | -34.7432 | -24.0219
14.815 | -34.7437 | -24.0298
14.8211 | -34.7441 | -24.0378
14.8272 | -34.7445 | -24.0461
14.8334 | -34.745 | -24.0539
14.8395 | -34.7455 | -24.0619
14.8456 | -34.7459 | -24.0699
14.8517 | -34.7464 | -24.0781
14.8578 | -34.7468 | -24.0859
14.8639 | -34.7473 | -24.0942
14.87 | -34.7478 | -24.1022
14.8761 | -34.7482 | -24.11
14.8822 | -34.7487 | -24.1181
14.8883 | -34.7491 | -24.126
14.8944 | -34.7496 | -24.1341
14.9005 | -34.7501 | -24.1419
14.9066 | -34.7505 | -24.1499
14.9127 | -34.751 | -24.1578
14.9188 | -34.7514 | -24.1661
14.9249 | -34.7519 | -24.174
14.931 | -34.7524 | -24.1819
14.9371 | -34.7528 | -24.19
14.9432 | -34.7533 | -24.1982
14.9493 | -34.7537 | -24.2061
14.9554 | -34.7541 | -24.214
14.9615 | -34.7545 | -24.2222
14.9676 | -34.755 | -24.2303
14.9738 | -34.7555 | -24.2382
14.9799 | -34.7559 | -24.2466
14.986 | -34.7564 | -24.2543
14.9921 | -34.7569 | -24.2622
14.9982 | -34.7573 | -24.2705
15.0043 | -34.7577 | -24.2785
15.0104 | -34.7582 | -24.2866
15.0165 | -34.7587 | -24.2948
15.0226 | -34.7591 | -24.3025
15.0287 | -34.7595 | -24.3109
15.0348 | -34.7601 | -24.3191
15.0409 | -34.7605 | -24.3274
15.047 | -34.761 | -24.3356
15.0531 | -34.7615 | -24.3438
15.0592 | -34.762 | -24.3519
15.0653 | -34.7625 | -24.3598
15.0714 | -34.7629 | -24.3674
15.0775 | -34.7634 | -24.3751
15.0836 | -34.7639 | -24.3832
15.0897 | -34.7644 | -24.3911
15.0958 | -34.7649 | -24.399
15.1019 | -34.7653 | -24.407
15.108 | -34.7658 | -24.415
15.1141 | -34.7662 | -24.423
15.1203 | -34.7667 | -24.4309
15.1264 | -34.7672 | -24.4392
15.1325 | -34.7677 | -24.447
15.1386 | -34.7682 | -24.455
15.1447 | -34.7687 | -24.463
15.1508 | -34.7692 | -24.4713
15.1569 | -34.7697 | -24.4796
15.163 | -34.7702 | -24.4873
15.1691 | -34.7707 | -24.4952
15.1752 | -34.7712 | -24.503
15.1813 | -34.7716 | -24.5111
15.1874 | -34.7721 | -24.5191
15.1935 | -34.7727 | -24.5268
15.1996 | -34.7732 | -24.5351
15.2057 | -34.7736 | -24.5429
15.2118 | -34.774 | -24.5509
15.2179 | -34.7745 | -24.5586
15.224 | -34.775 | -24.5667
15.2301 | -34.7755 | -24.5745
15.2362 | -34.7759 | -24.5824
15.2423 | -34.7764 | -24.5906
15.2484 | -34.7769 | -24.5984
15.2545 | -34.7774 | -24.6063
15.2607 | -34.7779 | -24.6144
15.2668 | -34.7785 | -24.6221
15.2729 | -34.7789 | -24.63
15.279 | -34.7794 | -24.6381
15.2851 | -34.7799 | -24.6462
15.2912 | -34.7804 | -24.6541
15.2973 | -34.7808 | -24.6621
15.3034 | -34.7814 | -24.6699
15.3095 | -34.7818 | -24.6777
15.3156 | -34.7823 | -24.6857
15.3217 | -34.7828 | -24.6933
15.3278 | -34.7833 | -24.7015
15.3339 | -34.7838 | -24.7091
15.34 | -34.7843 | -24.7167
15.3461 | -34.7848 | -24.7248
15.3522 | -34.7852 | -24.7326
15.3583 | -34.7857 | -24.7406
15.3644 | -34.7862 | -24.7486
15.3705 | -34.7867 | -24.7563
15.3766 | -34.7872 | -24.7647
15.3827 | -34.7876 | -24.7723
15.3888 | -34.7881 | -24.7799
15.3949 | -34.7886 | -24.7876
15.401 | -34.7891 | -24.7958
15.4072 | -34.7896 | -24.8034
15.4133 | -34.7901 | -24.8114
15.4194 | -34.7905 | -24.8194
15.4255 | -34.791 | -24.8273
15.4316 | -34.7915 | -24.8353
15.4377 | -34.792 | -24.843
15.4438 | -34.7925 | -24.8507
15.4499 | -34.793 | -24.8586
15.456 | -34.7934 | -24.8665
15.4621 | -34.794 | -24.8744
15.4682 | -34.7945 | -24.8825
15.4743 | -34.795 | -24.8902
15.4804 | -34.7955 | -24.8982
15.4865 | -34.7959 | -24.9061
15.4926 | -34.7964 | -24.9142
15.4987 | -34.7969 | -24.9223
15.5048 | -34.7974 | -24.93
15.5109 | -34.798 | -24.9377
15.517 | -34.7985 | -24.9457
15.5231 | -34.799 | -24.9536
15.5292 | -34.7994 | -24.9613
15.5353 | -34.7999 | -24.969
15.5414 | -34.8004 | -24.9771
15.5476 | -34.8009 | -24.9852
15.5537 | -34.8014 | -24.9931
15.5598 | -34.8019 | -25.0006
15.5659 | -34.8024 | -25.0088
15.572 | -34.8029 | -25.0163
15.5781 | -34.8034 | -25.0242
15.5842 | -34.8039 | -25.0319
15.5903 | -34.8044 | -25.0394
15.5964 | -34.805 | -25.047
15.6025 | -34.8054 | -25.0547
15.6086 | -34.8059 | -25.0626
15.6147 | -34.8064 | -25.0704
15.6208 | -34.807 | -25.0781
15.6269 | -34.8075 | -25.0858
15.633 | -34.808 | -25.0937
15.6391 | -34.8084 | -25.1016
15.6452 | -34.8089 | -25.1093
15.6513 | -34.8095 | -25.1174
15.6574 | -34.8099 | -25.1251
15.6635 | -34.8104 | -25.133
15.6696 | -34.811 | -25.1407
15.6757 | -34.8115 | -25.1488
15.6818 | -34.812 | -25.1569
15.688 | -34.8125 | -25.1648
15.6941 | -34.813 | -25.1727
15.7002 | -34.8135 | -25.1805
15.7063 | -34.814 | -25.1885
15.7124 | -34.8145 | -25.1964
15.7185 | -34.815 | -25.2042
15.7246 | -34.8155 | -25.2119
15.7307 | -34.816 | -25.2199
15.7368 | -34.8165 | -25.228
15.7429 | -34.817 | -25.2358
15.749 | -34.8175 | -25.2437
15.7551 | -34.818 | -25.2515
15.7612 | -34.8185 | -25.2588
15.7673 | -34.8189 | -25.2665
15.7734 | -34.8194 | -25.2742
15.7795 | -34.8199 | -25.2822
15.7856 | -34.8204 | -25.2898
15.7917 | -34.8209 | -25.2974
15.7978 | -34.8215 | -25.3053
15.8039 | -34.8219 | -25.3129
15.81 | -34.8224 | -25.3205
15.8161 | -34.8229 | -25.3285
15.8222 | -34.8234 | -25.3364
15.8283 | -34.8239 | -25.3442
15.8345 | -34.8244 | -25.352
15.8406 | -34.8249 | -25.3597
15.8467 | -34.8255 | -25.3675
15.8528 | -34.826 | -25.3754
15.8589 | -34.8265 | -25.3831
15.865 | -34.827 | -25.3907
15.8711 | -34.8275 | -25.3989
15.8772 | -34.828 | -25.4066
15.8833 | -34.8285 | -25.4149
15.8894 | -34.829 | -25.4226
15.8955 | -34.8295 | -25.4305
15.9016 | -34.8301 | -25.4381
15.9077 | -34.8306 | -25.4459
15.9138 | -34.8311 | -25.4535
15.9199 | -34.8316 | -25.4615
15.926 | -34.8321 | -25.4695
15.9321 | -34.8326 | -25.4773
15.9382 | -34.8332 | -25.485
15.9443 | -34.8337 | -25.493
15.9504 | -34.8342 | -25.5008
15.9565 | -34.8348 | -25.5086
15.9626 | -34.8353 | -25.5165
15.9687 | -34.8358 | -25.5242
15.9749 | -34.8364 | -25.532
15.981 | -34.8369 | -25.5397
15.9871 | -34.8374 | -25.5475
15.9932 | -34.8379 | -25.5553
15.9993 | -34.8384 | -25.5631
16.0054 | -34.8389 | -25.5706
16.0115 | -34.8394 | -25.5782
16.0176 | -34.8399 | -25.5863
16.0237 | -34.8405 | -25.5944
16.0298 | -34.841 | -25.6019
16.0359 | -34.8415 | -25.6094
16.042 | -34.842 | -25.6173
16.0481 | -34.8425 | -25.6251
16.0542 | -34.8431 | -25.6332
16.0603 | -34.8436 | -25.6407
16.0664 | -34.8441 | -25.6486
16.0725 | -34.8446 | -25.6562
16.0786 | -34.8452 | -25.664
16.0847 | -34.8456 | -25.6715
16.0908 | -34.8461 | -25.6791
16.0969 | -34.8466 | -25.6866
16.103 | -34.8471 | -25.6946
16.1091 | -34.8477 | -25.7023
16.1152 | -34.8482 | -25.7099
16.1214 | -34.8487 | -25.7178
16.1275 | -34.8492 | -25.7254
16.1336 | -34.8498 | -25.733
16.1397 | -34.8503 | -25.7407
16.1458 | -34.8508 | -25.7485
16.1519 | -34.8513 | -25.7564
16.158 | -34.8518 | -25.7644
16.1641 | -34.8523 | -25.7722
16.1702 | -34.8528 | -25.7802
16.1763 | -34.8533 | -25.788
16.1824 | -34.8539 | -25.7954
16.1885 | -34.8544 | -25.8032
16.1946 | -34.8549 | -25.8109
16.2007 | -34.8554 | -25.819
16.2068 | -34.8559 | -25.8265
16.2129 | -34.8564 | -25.8343
16.219 | -34.8569 | -25.8419
16.2251 | -34.8574 | -25.85
16.2312 | -34.8579 | -25.858
16.2373 | -34.8584 | -25.8657
16.2434 | -34.8589 | -25.8733
16.2495 | -34.8595 | -25.8809
16.2556 | -34.86 | -25.8888
16.2618 | -34.8605 | -25.8962
16.2679 | -34.861 | -25.9041
16.274 | -34.8616 | -25.9121
16.2801 | -34.8621 | -25.9195
16.2862 | -34.8626 | -25.9271
16.2923 | -34.8632 | -25.935
16.2984 | -34.8637 | -25.9429
16.3045 | -34.8643 | -25.9504
16.3106 | -34.8648 | -25.9577
16.3167 | -34.8653 | -25.9656
16.3228 | -34.8658 | -25.9735
16.3289 | -34.8663 | -25.9812
16.335 | -34.8668 | -25.9892
16.3411 | -34.8673 | -25.9965
16.3472 | -34.8678 | -26.0045
16.3533 | -34.8683 | -26.0123
16.3594 | -34.8688 | -26.02
16.3655 | -34.8693 | -26.028
16.3716 | -34.8698 | -26.0355
16.3777 | -34.8703 | -26.0433
16.3838 | -34.8708 | -26.0509
16.3899 | -34.8713 | -26.0589
16.396 | -34.8718 | -26.0668
16.4021 | -34.8723 | -26.0744
16.4083 | -34.8729 | -26.082
16.4144 | -34.8734 | -26.0896
16.4205 | -34.874 | -26.0976
16.4266 | -34.8744 | -26.105
16.4327 | -34.875 | -26.1123
16.4388 | -34.8755 | -26.1201
16.4449 | -34.876 | -26.1279
16.451 | -34.8765 | -26.1358
16.4571 | -34.877 | -26.1438
16.4632 | -34.8774 | -26.1514
16.4693 | -34.8779 | -26.1592
16.4754 | -34.8785 | -26.167
16.4815 | -34.879 | -26.1749
16.4876 | -34.8795 | -26.1826
16.4937 | -34.88 | -26.1901
16.4998 | -34.8805 | -26.1979
16.5059 | -34.881 | -26.2056
16.512 | -34.8815 | -26.2136
16.5181 | -34.8821 | -26.2216
16.5242 | -34.8826 | -26.2294
16.5303 | -34.8831 | -26.2368
16.5364 | -34.8836 | -26.2444
16.5425 | -34.8841 | -26.2522
16.5487 | -34.8846 | -26.2599
16.5548 | -34.8851 | -26.2676
16.5609 | -34.8856 | -26.2756
16.567 | -34.8861 | -26.2835
16.5731 | -34.8866 | -26.2914
16.5792 | -34.8871 | -26.299
16.5853 | -34.8877 | -26.3068
16.5914 | -34.8882 | -26.3147
16.5975 | -34.8887 | -26.3227
16.6036 | -34.8892 | -26.3303
16.6097 | -34.8898 | -26.3379
16.6158 | -34.8902 | -26.3457
16.6219 | -34.8908 | -26.3533
16.628 | -34.8913 | -26.3611
16.6341 | -34.8918 | -26.3686
16.6402 | -34.8923 | -26.3757
16.6463 | -34.8928 | -26.3833
16.6524 | -34.8933 | -26.3909
16.6585 | -34.8939 | -26.3989
16.6646 | -34.8944 | -26.4068
16.6707 | -34.8949 | -26.4142
16.6768 | -34.8954 | -26.4219
16.6829 | -34.896 | -26.4296
16.689 | -34.8965 | -26.4375
16.6952 | -34.897 | -26.4453
16.7013 | -34.8975 | -26.4531
16.7074 | -34.898 | -26.4605
16.7135 | -34.8985 | -26.4682
16.7196 | -34.899 | -26.4759
16.7257 | -34.8996 | -26.4835
16.7318 | -34.9 | -26.4912
16.7379 | -34.9006 | -26.4991
16.744 | -34.9011 | -26.5063
16.7501 | -34.9016 | -26.5142
16.7562 | -34.9021 | -26.5218
16.7623 | -34.9026 | -26.5295
16.7684 | -34.9032 | -26.5372
16.7745 | -34.9037 | -26.5446
16.7806 | -34.9042 | -26.5524
16.7867 | -34.9047 | -26.56
16.7928 | -34.9052 | -26.5677
16.7989 | -34.9057 | -26.5754
16.805 | -34.9062 | -26.5831
16.8111 | -34.9067 | -26.5907
16.8172 | -34.9072 | -26.5988
16.8233 | -34.9077 | -26.6067
16.8294 | -34.9083 | -26.6146
16.8356 | -34.9088 | -26.6224
16.8417 | -34.9093 | -26.6301
16.8478 | -34.9098 | -26.6379
16.8539 | -34.9103 | -26.6455
16.86 | -34.9108 | -26.6532
16.8661 | -34.9113 | -26.6609
16.8722 | -34.9118 | -26.6685
16.8783 | -34.9124 | -26.6759
16.8844 | -34.9129 | -26.6836
16.8905 | -34.9134 | -26.6914
16.8966 | -34.9139 | -26.6992
16.9027 | -34.9144 | -26.7064
16.9088 | -34.9149 | -26.7144
16.9149 | -34.9154 | -26.7224
16.921 | -34.916 | -26.7299
16.9271 | -34.9165 | -26.7379
16.9332 | -34.917 | -26.7458
16.9393 | -34.9175 | -26.7536
16.9454 | -34.9179 | -26.761
16.9515 | -34.9184 | -26.7689
16.9576 | -34.9189 | -26.7767
16.9637 | -34.9195 | -26.7842
16.9698 | -34.92 | -26.7917
16.9759 | -34.9205 | -26.7993
16.9821 | -34.921 | -26.8073
16.9882 | -34.9215 | -26.8149
16.9943 | -34.922 | -26.8227
17.0004 | -34.9225 | -26.8304
17.0065 | -34.923 | -26.8381
17.0126 | -34.9235 | -26.8458
17.0187 | -34.924 | -26.8537
17.0248 | -34.9245 | -26.8612
17.0309 | -34.925 | -26.8688
17.037 | -34.9255 | -26.8767
17.0431 | -34.9261 | -26.8841
17.0492 | -34.9265 | -26.8916
17.0553 | -34.927 | -26.8993
17.0614 | -34.9275 | -26.907
17.0675 | -34.928 | -26.9147
17.0736 | -34.9285 | -26.9225
17.0797 | -34.929 | -26.9297
17.0858 | -34.9295 | -26.9377
17.0919 | -34.93 | -26.9454
17.098 | -34.9305 | -26.9529
17.1041 | -34.931 | -26.9604
17.1102 | -34.9315 | -26.9683
17.1163 | -34.932 | -26.9761
17.1225 | -34.9325 | -26.9833
17.1286 | -34.9331 | -26.9907
17.1347 | -34.9335 | -26.9983
17.1408 | -34.9341 | -27.0059
17.1469 | -34.9346 | -27.0134
17.153 | -34.9351 | -27.0211
17.1591 | -34.9356 | -27.0285
17.1652 | -34.9361 | -27.0359
17.1713 | -34.9366 | -27.0436
17.1774 | -34.9371 | -27.0516
17.1835 | -34.9376 | -27.0594
17.1896 | -34.938 | -27.0668
17.1957 | -34.9385 | -27.0746
17.2018 | -34.939 | -27.0822
17.2079 | -34.9395 | -27.0896
17.214 | -34.9399 | -27.0971
17.2201 | -34.9404 | -27.1047
17.2262 | -34.9409 | -27.1124
17.2323 | -34.9414 | -27.1199
17.2384 | -34.942 | -27.1274
17.2445 | -34.9425 | -27.1349
17.2506 | -34.943 | -27.1425
17.2567 | -34.9434 | -27.1502
17.2628 | -34.9439 | -27.1583
17.269 | -34.9444 | -27.1661
17.2751 | -34.9449 | -27.1737
17.2812 | -34.9454 | -27.1812
17.2873 | -34.9459 | -27.189
17.2934 | -34.9464 | -27.1965
17.2995 | -34.9468 | -27.2039
17.3056 | -34.9473 | -27.2114
17.3117 | -34.9478 | -27.2189
17.3178 | -34.9483 | -27.2265
17.3239 | -34.9488 | -27.2339
17.33 | -34.9493 | -27.2413
17.3361 | -34.9497 | -27.2488
17.3422 | -34.9502 | -27.2566
17.3483 | -34.9507 | -27.2642
17.3544 | -34.9512 | -27.2718
17.3605 | -34.9517 | -27.2792
17.3666 | -34.9522 | -27.287
17.3727 | -34.9527 | -27.2946
17.3788 | -34.9532 | -27.302
17.3849 | -34.9537 | -27.3096
17.391 | -34.9542 | -27.317
17.3971 | -34.9547 | -27.3245
17.4032 | -34.9552 | -27.3325
17.4094 | -34.9557 | -27.34
17.4155 | -34.9562 | -27.3474
17.4216 | -34.9566 | -27.355
17.4277 | -34.9571 | -27.3628
17.4338 | -34.9577 | -27.3704
17.4399 | -34.9581 | -27.378
17.446 | -34.9586 | -27.3852
17.4521 | -34.9591 | -27.3927
17.4582 | -34.9596 | -27.4001
17.4643 | -34.9601 | -27.4078
17.4704 | -34.9605 | -27.4153
17.4765 | -34.961 | -27.423
17.4826 | -34.9615 | -27.4305
17.4887 | -34.962 | -27.4382
17.4948 | -34.9625 | -27.4457
17.5009 | -34.9629 | -27.4531
17.507 | -34.9635 | -27.4608
17.5131 | -34.9639 | -27.4684
17.5192 | -34.9644 | -27.476
17.5253 | -34.9649 | -27.4835
17.5314 | -34.9653 | -27.491
17.5375 | -34.9658 | -27.4988
17.5436 | -34.9663 | -27.5064
17.5497 | -34.9668 | -27.5142
17.5559 | -34.9672 | -27.5215
17.562 | -34.9677 | -27.5288
17.5681 | -34.9682 | -27.5363
17.5742 | -34.9687 | -27.5439
17.5803 | -34.9692 | -27.5514
17.5864 | -34.9697 | -27.5587
17.5925 | -34.9702 | -27.5666
17.5986 | -34.9707 | -27.5738
17.6047 | -34.9712 | -27.5816
17.6108 | -34.9717 | -27.5889
17.6169 | -34.9722 | -27.5965
17.623 | -34.9727 | -27.6039
17.6291 | -34.9732 | -27.6114
17.6352 | -34.9737 | -27.6188
17.6413 | -34.9741 | -27.6265
17.6474 | -34.9746 | -27.6341
17.6535 | -34.9751 | -27.6416
17.6596 | -34.9756 | -27.649
17.6657 | -34.9761 | -27.6564
17.6718 | -34.9765 | -27.664
17.6779 | -34.977 | -27.6713
17.684 | -34.9775 | -27.6789
17.6901 | -34.9781 | -27.6864
17.6963 | -34.9785 | -27.6937
17.7024 | -34.979 | -27.7011
17.7085 | -34.9794 | -27.7084
17.7146 | -34.98 | -27.7159
17.7207 | -34.9804 | -27.7234
17.7268 | -34.9809 | -27.7308
17.7329 | -34.9814 | -27.7386
17.739 | -34.9819 | -27.7462
17.7451 | -34.9824 | -27.7536
17.7512 | -34.9829 | -27.7612
17.7573 | -34.9834 | -27.7686
17.7634 | -34.9839 | -27.7757
17.7695 | -34.9843 | -27.7833
17.7756 | -34.9849 | -27.7907
17.7817 | -34.9854 | -27.7981
17.7878 | -34.9858 | -27.8056
17.7939 | -34.9863 | -27.813
17.8 | -34.9868 | -27.8203
17.8061 | -34.9873 | -27.8279
17.8122 | -34.9878 | -27.8355
17.8183 | -34.9883 | -27.843
17.8244 | -34.9888 | -27.8504
17.8305 | -34.9893 | -27.8578
17.8366 | -34.9897 | -27.8649
17.8428 | -34.9902 | -27.8726
17.8489 | -34.9907 | -27.8798
17.855 | -34.9912 | -27.887
17.8611 | -34.9917 | -27.8944
17.8672 | -34.9922 | -27.9015
17.8733 | -34.9926 | -27.9092
17.8794 | -34.9931 | -27.9169
17.8855 | -34.9936 | -27.9242
17.8916 | -34.9941 | -27.9315
17.8977 | -34.9946 | -27.939
17.9038 | -34.9951 | -27.9462
17.9099 | -34.9955 | -27.9537
17.916 | -34.996 | -27.9611
17.9221 | -34.9965 | -27.9685
17.9282 | -34.997 | -27.9759
17.9343 | -34.9975 | -27.9834
17.9404 | -34.9979 | -27.9907
17.9465 | -34.9984 | -27.9983
17.9526 | -34.9988 | -28.0059
17.9587 | -34.9993 | -28.0133
17.9648 | -34.9998 | -28.0207
17.9709 | -35.0003 | -28.0282
17.977 | -35.0008 | -28.0355
17.9832 | -35.0013 | -28.043
17.9893 | -35.0017 | -28.0506
17.9954 | -35.0022 | -28.0582
18.0015 | -35.0027 | -28.0658
18.0076 | -35.0032 | -28.0728
18.0137 | -35.0037 | -28.0805
18.0198 | -35.0041 | -28.088
18.0259 | -35.0046 | -28.0956
18.032 | -35.005 | -28.1027
18.0381 | -35.0055 | -28.11
18.0442 | -35.006 | -28.1176
18.0503 | -35.0065 | -28.1251
18.0564 | -35.007 | -28.1324
18.0625 | -35.0074 | -28.1399
18.0686 | -35.0079 | -28.1472
18.0747 | -35.0084 | -28.1545
18.0808 | -35.0089 | -28.1616
18.0869 | -35.0094 | -28.1691
18.093 | -35.0099 | -28.1765
18.0991 | -35.0104 | -28.1842
18.1052 | -35.0109 | -28.1918
18.1113 | -35.0114 | -28.1987
18.1174 | -35.0118 | -28.2059
18.1236 | -35.0123 | -28.2134
18.1297 | -35.0128 | -28.2205
18.1358 | -35.0133 | -28.228
18.1419 | -35.0139 | -28.2354
18.148 | -35.0144 | -28.2429
18.1541 | -35.0149 | -28.2498
18.1602 | -35.0154 | -28.2574
18.1663 | -35.0159 | -28.2648
18.1724 | -35.0164 | -28.2724
18.1785 | -35.0169 | -28.2799
18.1846 | -35.0173 | -28.2871
18.1907 | -35.0177 | -28.2946
18.1968 | -35.0182 | -28.3021
18.2029 | -35.0187 | -28.3092
18.209 | -35.0192 | -28.3169
18.2151 | -35.0197 | -28.3245
18.2212 | -35.0202 | -28.3316
18.2273 | -35.0208 | -28.339
18.2334 | -35.0212 | -28.3464
18.2395 | -35.0217 | -28.3536
18.2456 | -35.0222 | -28.361
18.2517 | -35.0226 | -28.3683
18.2578 | -35.0231 | -28.3754
18.2639 | -35.0236 | -28.3829
18.2701 | -35.0241 | -28.3902
18.2762 | -35.0246 | -28.3977
18.2823 | -35.0251 | -28.4051
18.2884 | -35.0256 | -28.4126
18.2945 | -35.026 | -28.4197
18.3006 | -35.0265 | -28.427
18.3067 | -35.027 | -28.4344
18.3128 | -35.0275 | -28.4418
18.3189 | -35.028 | -28.4493
18.325 | -35.0285 | -28.4563
18.3311 | -35.029 | -28.4638
18.3372 | -35.0295 | -28.4708
18.3433 | -35.0299 | -28.478
18.3494 | -35.0304 | -28.4852
18.3555 | -35.0309 | -28.4922
18.3616 | -35.0314 | -28.4994
18.3677 | -35.0319 | -28.5067
18.3738 | -35.0324 | -28.5139
18.3799 | -35.0328 | -28.5211
18.386 | -35.0333 | -28.5284
18.3921 | -35.0338 | -28.5355
18.3982 | -35.0342 | -28.5425
18.4043 | -35.0347 | -28.5497
18.4105 | -35.0351 | -28.5568
18.4166 | -35.0356 | -28.5641
18.4227 | -35.0361 | -28.5715
18.4288 | -35.0366 | -28.579
18.4349 | -35.0371 | -28.5863
18.441 | -35.0376 | -28.5935
18.4471 | -35.0381 | -28.601
18.4532 | -35.0385 | -28.6084
18.4593 | -35.039 | -28.6155
18.4654 | -35.0395 | -28.6226
18.4715 | -35.04 | -28.63
18.4776 | -35.0405 | -28.6376
18.4837 | -35.041 | -28.645
18.4898 | -35.0415 | -28.6525
18.4959 | -35.042 | -28.6598
18.502 | -35.0425 | -28.6671
18.5081 | -35.043 | -28.6747
18.5142 | -35.0434 | -28.6823
18.5203 | -35.0439 | -28.6892
18.5264 | -35.0444 | -28.6962
18.5325 | -35.0449 | -28.7036
18.5386 | -35.0454 | -28.7112
18.5447 | -35.0459 | -28.7184
18.5508 | -35.0464 | -28.7256
18.557 | -35.0469 | -28.7331
18.5631 | -35.0473 | -28.7403
18.5692 | -35.0478 | -28.7474
18.5753 | -35.0483 | -28.7544
18.5814 | -35.0488 | -28.7619
18.5875 | -35.0492 | -28.7695
18.5936 | -35.0497 | -28.7772
18.5997 | -35.0502 | -28.7846
18.6058 | -35.0507 | -28.7918
18.6119 | -35.0512 | -28.7992
18.618 | -35.0517 | -28.8065
18.6241 | -35.0522 | -28.814
18.6302 | -35.0527 | -28.8217
18.6363 | -35.0532 | -28.8285
18.6424 | -35.0537 | -28.8357
18.6485 | -35.0541 | -28.8431
18.6546 | -35.0546 | -28.8502
18.6607 | -35.055 | -28.8573
18.6668 | -35.0556 | -28.8646
18.6729 | -35.0561 | -28.872
18.679 | -35.0566 | -28.879
18.6851 | -35.057 | -28.8862
18.6912 | -35.0575 | -28.8936
18.6974 | -35.058 | -28.901
18.7035 | -35.0585 | -28.9083
18.7096 | -35.059 | -28.9157
18.7157 | -35.0595 | -28.9229
18.7218 | -35.06 | -28.9303
18.7279 | -35.0605 | -28.9373
18.734 | -35.061 | -28.9445
18.7401 | -35.0615 | -28.9515
18.7462 | -35.062 | -28.9581
18.7523 | -35.0625 | -28.9653
18.7584 | -35.063 | -28.9728
18.7645 | -35.0635 | -28.9799
18.7706 | -35.064 | -28.9872
18.7767 | -35.0645 | -28.9945
18.7828 | -35.065 | -29.0013
18.7889 | -35.0654 | -29.0084
18.795 | -35.0659 | -29.0156
18.8011 | -35.0664 | -29.023
18.8072 | -35.0669 | -29.0303
18.8133 | -35.0673 | -29.0376
18.8194 | -35.0678 | -29.0446
18.8255 | -35.0683 | -29.0518
18.8316 | -35.0688 | -29.059
18.8377 | -35.0692 | -29.066
18.8439 | -35.0697 | -29.073
18.85 | -35.0702 | -29.0803
18.8561 | -35.0707 | -29.0875
18.8622 | -35.0711 | -29.0949
18.8683 | -35.0716 | -29.1019
18.8744 | -35.0721 | -29.1089
18.8805 | -35.0726 | -29.1162
18.8866 | -35.0731 | -29.1235
18.8927 | -35.0737 | -29.1304
18.8988 | -35.0741 | -29.1376
18.9049 | -35.0746 | -29.1447
18.911 | -35.0752 | -29.1517
18.9171 | -35.0757 | -29.1589
18.9232 | -35.0762 | -29.1659
18.9293 | -35.0767 | -29.1732
18.9354 | -35.0771 | -29.1801
18.9415 | -35.0776 | -29.187
18.9476 | -35.078 | -29.1942
18.9537 | -35.0785 | -29.2017
18.9598 | -35.0791 | -29.2083
18.9659 | -35.0796 | -29.2155
18.972 | -35.08 | -29.2226
18.9781 | -35.0805 | -29.2296
18.9843 | -35.081 | -29.2371
18.9904 | -35.0814 | -29.2443
18.9965 | -35.0819 | -29.2515
19.0026 | -35.0824 | -29.2587
19.0087 | -35.0829 | -29.2659
19.0148 | -35.0834 | -29.2731
19.0209 | -35.0839 | -29.2805
19.027 | -35.0844 | -29.2879
19.0331 | -35.0849 | -29.2952
19.0392 | -35.0854 | -29.3024
19.0453 | -35.0859 | -29.3096
19.0514 | -35.0864 | -29.3168
19.0575 | -35.0868 | -29.324
19.0636 | -35.0874 | -29.3306
19.0697 | -35.0879 | -29.338
19.0758 | -35.0884 | -29.345
19.0819 | -35.0889 | -29.3523
19.088 | -35.0894 | -29.3592
19.0941 | -35.0899 | -29.3663
19.1002 | -35.0904 | -29.3733
19.1063 | -35.0909 | -29.3805
19.1124 | -35.0914 | -29.3876
19.1185 | -35.0919 | -29.3949
19.1246 | -35.0924 | -29.402
19.1308 | -35.0929 | -29.4089
19.1369 | -35.0934 | -29.4157
19.143 | -35.0939 | -29.4231
19.1491 | -35.0944 | -29.4303
19.1552 | -35.0949 | -29.4377
19.1613 | -35.0954 | -29.4447
19.1674 | -35.0959 | -29.452
19.1735 | -35.0964 | -29.4592
19.1796 | -35.0969 | -29.4662
19.1857 | -35.0974 | -29.4734
19.1918 | -35.0979 | -29.4807
19.1979 | -35.0984 | -29.4882
19.204 | -35.0989 | -29.4953
19.2101 | -35.0994 | -29.5023
19.2162 | -35.0999 | -29.5097
19.2223 | -35.1003 | -29.5166
19.2284 | -35.1008 | -29.5238
19.2345 | -35.1013 | -29.5312
19.2406 | -35.1018 | -29.5387
19.2467 | -35.1022 | -29.5458
19.2528 | -35.1028 | -29.5529
19.2589 | -35.1032 | -29.56
19.265 | -35.1037 | -29.5671
19.2712 | -35.1042 | -29.5742
19.2773 | -35.1047 | -29.5811
19.2834 | -35.1052 | -29.588
19.2895 | -35.1057 | -29.5953
19.2956 | -35.1062 | -29.6026
19.3017 | -35.1067 | -29.6096
19.3078 | -35.1072 | -29.6166
19.3139 | -35.1077 | -29.6239
19.32 | -35.1082 | -29.6311
19.3261 | -35.1086 | -29.6382
19.3322 | -35.1092 | -29.6452
19.3383 | -35.1097 | -29.6522
19.3444 | -35.1102 | -29.6591
19.3505 | -35.1107 | -29.6661
19.3566 | -35.1112 | -29.6733
19.3627 | -35.1117 | -29.6804
19.3688 | -35.1121 | -29.6875
19.3749 | -35.1126 | -29.6948
19.381 | -35.1132 | -29.7017
19.3871 | -35.1136 | -29.7088
19.3932 | -35.1141 | -29.7157
19.3993 | -35.1146 | -29.7228
19.4054 | -35.1151 | -29.7302
19.4115 | -35.1156 | -29.7374
19.4177 | -35.116 | -29.7444
19.4238 | -35.1165 | -29.7513
19.4299 | -35.1171 | -29.7583
19.436 | -35.1176 | -29.7659
19.4421 | -35.1181 | -29.7724
19.4482 | -35.1186 | -29.7796
19.4543 | -35.1191 | -29.7862
19.4604 | -35.1196 | -29.7931
19.4665 | -35.1201 | -29.8001
19.4726 | -35.1206 | -29.807
19.4787 | -35.1211 | -29.8141
19.4848 | -35.1216 | -29.821
19.4909 | -35.1221 | -29.828
19.497 | -35.1225 | -29.835
19.5031 | -35.123 | -29.842
19.5092 | -35.1235 | -29.8489
19.5153 | -35.124 | -29.8561
19.5214 | -35.1245 | -29.863
19.5275 | -35.125 | -29.8699
19.5336 | -35.1254 | -29.8768
19.5397 | -35.126 | -29.8837
19.5458 | -35.1265 | -29.8907
19.5519 | -35.127 | -29.8976
19.5581 | -35.1275 | -29.9044
19.5642 | -35.128 | -29.9116
19.5703 | -35.1285 | -29.9184
19.5764 | -35.129 | -29.9255
19.5825 | -35.1295 | -29.9326
19.5886 | -35.13 | -29.9396
19.5947 | -35.1305 | -29.9468
19.6008 | -35.131 | -29.9539
19.6069 | -35.1315 | -29.9609
19.613 | -35.132 | -29.9681
19.6191 | -35.1325 | -29.975
19.6252 | -35.133 | -29.982
19.6313 | -35.1335 | -29.9888
19.6374 | -35.134 | -29.996
19.6435 | -35.1345 | -30.003
19.6496 | -35.135 | -30.01
19.6557 | -35.1355 | -30.017
19.6618 | -35.136 | -30.0236
19.6679 | -35.1365 | -30.0307
19.674 | -35.137 | -30.0375
19.6801 | -35.1375 | -30.0446
19.6862 | -35.138 | -30.0516
19.6923 | -35.1385 | -30.0586
19.6984 | -35.139 | -30.0653
19.7046 | -35.1396 | -30.0724
19.7107 | -35.14 | -30.0794
19.7168 | -35.1405 | -30.0867
19.7229 | -35.141 | -30.0938
19.729 | -35.1414 | -30.101
19.7351 | -35.1419 | -30.1075
19.7412 | -35.1425 | -30.1143
19.7473 | -35.1429 | -30.1212
19.7534 | -35.1435 | -30.1283
19.7595 | -35.144 | -30.1354
19.7656 | -35.1445 | -30.1425
19.7717 | -35.1449 | -30.1493
19.7778 | -35.1454 | -30.1563
19.7839 | -35.1459 | -30.1632
19.79 | -35.1464 | -30.17
19.7961 | -35.1469 | -30.1773
19.8022 | -35.1474 | -30.1843
19.8083 | -35.1479 | -30.1913
19.8144 | -35.1484 | -30.1983
19.8205 | -35.1489 | -30.2054
19.8266 | -35.1494 | -30.2122
19.8327 | -35.1499 | -30.2193
19.8388 | -35.1504 | -30.2261
19.845 | -35.1509 | -30.2332
19.8511 | -35.1514 | -30.2404
19.8572 | -35.1518 | -30.2477
19.8633 | -35.1523 | -30.2546
19.8694 | -35.1529 | -30.2617
19.8755 | -35.1534 | -30.2687
19.8816 | -35.154 | -30.2755
19.8877 | -35.1545 | -30.2822
19.8938 | -35.1549 | -30.2891
19.8999 | -35.1554 | -30.2961
19.906 | -35.1559 | -30.303
19.9121 | -35.1564 | -30.3101
19.9182 | -35.1569 | -30.3165
19.9243 | -35.1574 | -30.324
19.9304 | -35.1578 | -30.3313
19.9365 | -35.1583 | -30.3384
19.9426 | -35.1588 | -30.345
19.9487 | -35.1593 | -30.3517
19.9548 | -35.1598 | -30.359
19.9609 | -35.1603 | -30.3659
19.967 | -35.1609 | -30.3727
19.9731 | -35.1614 | -30.3795
19.9792 | -35.1618 | -30.3865
19.9853 | -35.1624 | -30.3935
19.9915 | -35.1629 | -30.4007
19.9976 | -35.1634 | -30.4075
20.0037 | -35.164 | -30.4144
20.0098 | -35.1644 | -30.4214
20.0159 | -35.1649 | -30.4283
20.022 | -35.1654 | -30.4356
20.0281 | -35.1659 | -30.4426
20.0342 | -35.1663 | -30.4493
20.0403 | -35.1669 | -30.456
20.0464 | -35.1673 | -30.4634
20.0525 | -35.1678 | -30.4698
20.0586 | -35.1683 | -30.4764
20.0647 | -35.1688 | -30.4836
20.0708 | -35.1693 | -30.4906
20.0769 | -35.1698 | -30.4976
20.083 | -35.1703 | -30.5041
20.0891 | -35.1707 | -30.511
20.0952 | -35.1712 | -30.5175
20.1013 | -35.1717 | -30.5244
20.1074 | -35.1721 | -30.5315
20.1135 | -35.1726 | -30.5382
20.1196 | -35.1731 | -30.5452
20.1257 | -35.1737 | -30.5521
20.1319 | -35.1742 | -30.559
20.138 | -35.1746 | -30.5659
20.1441 | -35.1751 | -30.5731
20.1502 | -35.1757 | -30.5803
20.1563 | -35.1761 | -30.5868
20.1624 | -35.1766 | -30.5938
20.1685 | -35.1771 | -30.601
20.1746 | -35.1776 | -30.6082
20.1807 | -35.1781 | -30.6151
20.1868 | -35.1786 | -30.6219
20.1929 | -35.1791 | -30.6283
20.199 | -35.1796 | -30.6356
20.2051 | -35.1801 | -30.6425
20.2112 | -35.1806 | -30.6491
20.2173 | -35.1811 | -30.656
20.2234 | -35.1816 | -30.6629
20.2295 | -35.1821 | -30.6693
20.2356 | -35.1826 | -30.6763
20.2417 | -35.1831 | -30.683
20.2478 | -35.1836 | -30.6895
20.2539 | -35.184 | -30.6965
20.26 | -35.1845 | -30.7031
20.2661 | -35.185 | -30.7102
20.2723 | -35.1855 | -30.7172
20.2784 | -35.1861 | -30.724
20.2845 | -35.1865 | -30.7309
20.2906 | -35.1871 | -30.7379
20.2967 | -35.1877 | -30.7447
20.3028 | -35.1882 | -30.7515
20.3089 | -35.1887 | -30.7583
20.315 | -35.1892 | -30.765
20.3211 | -35.1896 | -30.7721
20.3272 | -35.1902 | -30.7789
20.3333 | -35.1907 | -30.7857
20.3394 | -35.1912 | -30.7924
20.3455 | -35.1917 | -30.7993
20.3516 | -35.1922 | -30.806
20.3577 | -35.1926 | -30.8127
20.3638 | -35.1932 | -30.8192
20.3699 | -35.1936 | -30.8258
20.376 | -35.1941 | -30.8326
20.3821 | -35.1946 | -30.8394
20.3882 | -35.1951 | -30.8462
20.3943 | -35.1957 | -30.8531
20.4004 | -35.1962 | -30.8595
20.4065 | -35.1967 | -30.8663
20.4126 | -35.1971 | -30.8727
20.4188 | -35.1977 | -30.8795
20.4249 | -35.1982 | -30.8867
20.431 | -35.1986 | -30.8934
20.4371 | -35.1992 | -30.8997
20.4432 | -35.1997 | -30.9068
20.4493 | -35.2002 | -30.9139
20.4554 | -35.2007 | -30.9207
20.4615 | -35.2011 | -30.9278
20.4676 | -35.2016 | -30.9345
20.4737 | -35.2021 | -30.9408
20.4798 | -35.2027 | -30.9476
20.4859 | -35.2032 | -30.954
20.492 | -35.2036 | -30.9609
20.4981 | -35.2041 | -30.9678
20.5042 | -35.2046 | -30.9743
20.5103 | -35.2051 | -30.9805
20.5164 | -35.2056 | -30.9874
20.5225 | -35.206 | -30.9945
20.5286 | -35.2065 | -31.0018
20.5347 | -35.207 | -31.0085
20.5408 | -35.2075 | -31.015
20.5469 | -35.208 | -31.0216
20.553 | -35.2086 | -31.0287
20.5592 | -35.209 | -31.0354
20.5653 | -35.2095 | -31.0423
20.5714 | -35.2099 | -31.0488
20.5775 | -35.2105 | -31.055
20.5836 | -35.2109 | -31.0621
20.5897 | -35.2115 | -31.0688
20.5958 | -35.2119 | -31.0755
20.6019 | -35.2124 | -31.0823
20.608 | -35.2128 | -31.0894
20.6141 | -35.2133 | -31.0958
20.6202 | -35.2138 | -31.103
20.6263 | -35.2143 | -31.1094
20.6324 | -35.2147 | -31.1157
20.6385 | -35.2152 | -31.1224
20.6446 | -35.2157 | -31.1293
20.6507 | -35.2162 | -31.1364
20.6568 | -35.2168 | -31.1429
20.6629 | -35.2172 | -31.1493
20.669 | -35.2177 | -31.1562
20.6751 | -35.2182 | -31.1629
20.6812 | -35.2187 | -31.169
20.6873 | -35.2191 | -31.1758
20.6934 | -35.2197 | -31.183
20.6995 | -35.2201 | -31.1895
20.7057 | -35.2206 | -31.1964
20.7118 | -35.2212 | -31.2028
20.7179 | -35.2217 | -31.2095
20.724 | -35.2222 | -31.216
20.7301 | -35.2226 | -31.2223
20.7362 | -35.2232 | -31.2293
20.7423 | -35.2236 | -31.2356
20.7484 | -35.2242 | -31.2417
20.7545 | -35.2247 | -31.2479
20.7606 | -35.2253 | -31.2552
20.7667 | -35.2258 | -31.2617
20.7728 | -35.2263 | -31.2682
20.7789 | -35.2268 | -31.2751
20.785 | -35.2273 | -31.2817
20.7911 | -35.2277 | -31.2879
20.7972 | -35.2282 | -31.2946
20.8033 | -35.2287 | -31.3013
20.8094 | -35.2292 | -31.3081
20.8155 | -35.2297 | -31.3149
20.8216 | -35.2303 | -31.322
20.8277 | -35.2307 | -31.3288
20.8338 | -35.2312 | -31.3351
20.8399 | -35.2318 | -31.3414
20.8461 | -35.2322 | -31.3477
20.8522 | -35.2327 | -31.3538
20.8583 | -35.2332 | -31.3612
20.8644 | -35.2337 | -31.3673
20.8705 | -35.2343 | -31.3742
20.8766 | -35.2348 | -31.3815
20.8827 | -35.2353 | -31.3885
20.8888 | -35.2357 | -31.3951
20.8949 | -35.2362 | -31.4025
20.901 | -35.2367 | -31.4088
20.9071 | -35.2373 | -31.4152
20.9132 | -35.2378 | -31.4217
20.9193 | -35.2382 | -31.4286
20.9254 | -35.2386 | -31.435
20.9315 | -35.2391 | -31.4409
20.9376 | -35.2395 | -31.4476
20.9437 | -35.24 | -31.4545
20.9498 | -35.2405 | -31.4614
20.9559 | -35.241 | -31.4678
20.962 | -35.2415 | -31.4743
20.9681 | -35.242 | -31.4814
20.9742 | -35.2425 | -31.488
20.9803 | -35.2429 | -31.4948
20.9864 | -35.2435 | -31.5014
20.9926 | -35.244 | -31.5086
20.9987 | -35.2443 | -31.5151
21.0048 | -35.2448 | -31.5213
21.0109 | -35.2452 | -31.5274
21.017 | -35.2457 | -31.5339
21.0231 | -35.2462 | -31.5407
21.0292 | -35.2468 | -31.5472
21.0353 | -35.2472 | -31.5541
21.0414 | -35.2477 | -31.5603
21.0475 | -35.2482 | -31.5668
21.0536 | -35.2487 | -31.5737
21.0597 | -35.2492 | -31.5804
21.0658 | -35.2496 | -31.5872
21.0719 | -35.2502 | -31.5935
21.078 | -35.2506 | -31.5997
21.0841 | -35.2511 | -31.6058
21.0902 | -35.2515 | -31.6119
21.0963 | -35.252 | -31.6183
21.1024 | -35.2526 | -31.6252
21.1085 | -35.2531 | -31.6306
21.1146 | -35.2535 | -31.6375
21.1207 | -35.254 | -31.6437
21.1268 | -35.2545 | -31.6501
21.133 | -35.2549 | -31.6575
21.1391 | -35.2554 | -31.6639
21.1452 | -35.2559 | -31.6713
21.1513 | -35.2564 | -31.6783
21.1574 | -35.2568 | -31.6844
21.1635 | -35.2574 | -31.691
21.1696 | -35.2578 | -31.6965
21.1757 | -35.2582 | -31.7038
21.1818 | -35.2587 | -31.7114
21.1879 | -35.2591 | -31.7178
21.194 | -35.2597 | -31.7246
21.2001 | -35.2601 | -31.7305
21.2062 | -35.2605 | -31.7366
21.2123 | -35.2609 | -31.7433
21.2184 | -35.2614 | -31.7494
21.2245 | -35.2618 | -31.7565
21.2306 | -35.2623 | -31.7633
21.2367 | -35.2628 | -31.7695
21.2428 | -35.2634 | -31.7758
21.2489 | -35.2638 | -31.7825
21.255 | -35.2643 | -31.7887
21.2611 | -35.2649 | -31.7948
21.2672 | -35.2653 | -31.8008
21.2733 | -35.2658 | -31.8077
21.2795 | -35.2661 | -31.8136
21.2856 | -35.2665 | -31.8204
21.2917 | -35.2669 | -31.8271
21.2978 | -35.2674 | -31.8334
21.3039 | -35.2678 | -31.8396
21.31 | -35.2683 | -31.8457
21.3161 | -35.2689 | -31.8515
21.3222 | -35.2694 | -31.8582
21.3283 | -35.2698 | -31.8647
21.3344 | -35.2703 | -31.8709
21.3405 | -35.2707 | -31.8777
21.3466 | -35.2712 | -31.8836
21.3527 | -35.2717 | -31.8899
21.3588 | -35.2723 | -31.8965
21.3649 | -35.2728 | -31.9029
21.371 | -35.2734 | -31.9096
21.3771 | -35.2738 | -31.9154
21.3832 | -35.2743 | -31.9215
21.3893 | -35.2747 | -31.9274
21.3954 | -35.2752 | -31.9341
21.4015 | -35.2757 | -31.9407
21.4076 | -35.2762 | -31.9474
21.4137 | -35.2767 | -31.9542
21.4199 | -35.2772 | -31.9607
21.426 | -35.2776 | -31.9672
21.4321 | -35.2781 | -31.9737
21.4382 | -35.2785 | -31.9804
21.4443 | -35.279 | -31.9871
21.4504 | -35.2795 | -31.9939
21.4565 | -35.2799 | -32.0001
21.4626 | -35.2803 | -32.0058
21.4687 | -35.2808 | -32.0125
21.4748 | -35.2811 | -32.0184
21.4809 | -35.2815 | -32.0249
21.487 | -35.2819 | -32.0315
21.4931 | -35.2825 | -32.0375
21.4992 | -35.283 | -32.0438
21.5053 | -35.2834 | -32.0501
21.5114 | -35.2839 | -32.0556
21.5175 | -35.2844 | -32.0619
21.5236 | -35.2849 | -32.0682
21.5297 | -35.2852 | -32.0747
21.5358 | -35.2855 | -32.0811
21.5419 | -35.2859 | -32.0871
21.548 | -35.2864 | -32.0937
21.5541 | -35.2869 | -32.1005
21.5602 | -35.2874 | -32.1073
21.5664 | -35.2879 | -32.1135
21.5725 | -35.2883 | -32.1203
21.5786 | -35.2888 | -32.1272
21.5847 | -35.2891 | -32.1339
21.5908 | -35.2895 | -32.1406
21.5969 | -35.2901 | -32.1471
21.603 | -35.2906 | -32.1528
21.6091 | -35.291 | -32.1595
21.6152 | -35.2915 | -32.1662
21.6213 | -35.2918 | -32.1728
21.6274 | -35.2923 | -32.1788
21.6335 | -35.2927 | -32.1845
21.6396 | -35.2932 | -32.1914
21.6457 | -35.2936 | -32.198
21.6518 | -35.294 | -32.2046
21.6579 | -35.2945 | -32.2111
21.664 | -35.2949 | -32.2185
21.6701 | -35.2952 | -32.2243
21.6762 | -35.2958 | -32.2302
21.6823 | -35.2962 | -32.236
21.6884 | -35.2967 | -32.2429
21.6945 | -35.2971 | -32.2493
21.7006 | -35.2975 | -32.2557
21.7068 | -35.298 | -32.2619
21.7129 | -35.2985 | -32.2684
21.719 | -35.299 | -32.2743
21.7251 | -35.2995 | -32.2803
21.7312 | -35.3 | -32.2871
21.7373 | -35.3005 | -32.2934
21.7434 | -35.3009 | -32.2996
21.7495 | -35.3013 | -32.306
21.7556 | -35.3017 | -32.3124
21.7617 | -35.3023 | -32.3186
21.7678 | -35.3027 | -32.3249
21.7739 | -35.303 | -32.3316
21.78 | -35.3034 | -32.3373
21.7861 | -35.3039 | -32.3441
21.7922 | -35.3044 | -32.3501
21.7983 | -35.3049 | -32.3564
21.8044 | -35.3054 | -32.3631
21.8105 | -35.3059 | -32.3694
21.8166 | -35.3064 | -32.3755
21.8227 | -35.3068 | -32.3816
21.8288 | -35.3072 | -32.3879
21.8349 | -35.3076 | -32.3937
21.841 | -35.3081 | -32.4
21.8471 | -35.3085 | -32.4065
21.8533 | -35.3089 | -32.4124
21.8594 | -35.3094 | -32.4194
21.8655 | -35.3098 | -32.4259
21.8716 | -35.3103 | -32.4327
21.8777 | -35.3106 | -32.4388
21.8838 | -35.3111 | -32.445
21.8899 | -35.3115 | -32.451
21.896 | -35.3121 | -32.457
21.9021 | -35.3126 | -32.4633
21.9082 | -35.313 | -32.4694
21.9143 | -35.3134 | -32.4752
21.9204 | -35.3138 | -32.4809
21.9265 | -35.3143 | -32.4866
21.9326 | -35.3148 | -32.4933
21.9387 | -35.3152 | -32.4994
21.9448 | -35.3157 | -32.5058
21.9509 | -35.3161 | -32.5122
21.957 | -35.3166 | -32.5186
21.9631 | -35.3172 | -32.5251
21.9692 | -35.3176 | -32.5319
21.9753 | -35.318 | -32.5379
21.9814 | -35.3185 | -32.5448
21.9875 | -35.3189 | -32.5513
21.9937 | -35.3194 | -32.5586
21.9998 | -35.3198 | -32.5654
22.0059 | -35.3202 | -32.5711
22.012 | -35.3206 | -32.5779
22.0181 | -35.3212 | -32.5847
22.0242 | -35.3217 | -32.591
22.0303 | -35.3221 | -32.5974
22.0364 | -35.3226 | -32.6042
22.0425 | -35.323 | -32.6106
22.0486 | -35.3234 | -32.6167
22.0547 | -35.3238 | -32.623
22.0608 | -35.3242 | -32.6292
22.0669 | -35.3246 | -32.6356
22.073 | -35.325 | -32.6416
22.0791 | -35.3254 | -32.6483
22.0852 | -35.3258 | -32.6547
22.0913 | -35.3263 | -32.6604
22.0974 | -35.3267 | -32.6664
22.1035 | -35.3271 | -32.6725
22.1096 | -35.3274 | -32.6788
22.1157 | -35.3279 | -32.6848
22.1218 | -35.3283 | -32.6904
22.1279 | -35.3288 | -32.6958
22.134 | -35.3292 | -32.7022
22.1402 | -35.3297 | -32.7086
22.1463 | -35.3302 | -32.715
22.1524 | -35.3307 | -32.7212
22.1585 | -35.3311 | -32.7276
22.1646 | -35.3316 | -32.7336
22.1707 | -35.3321 | -32.7407
22.1768 | -35.3325 | -32.7471
22.1829 | -35.3329 | -32.7541
22.189 | -35.3333 | -32.7602
22.1951 | -35.3338 | -32.7665
22.2012 | -35.3342 | -32.7728
22.2073 | -35.3346 | -32.7794
22.2134 | -35.3351 | -32.7857
22.2195 | -35.3356 | -32.7922
22.2256 | -35.336 | -32.7986
22.2317 | -35.3366 | -32.8044
22.2378 | -35.337 | -32.8102
22.2439 | -35.3375 | -32.8164
22.25 | -35.338 | -32.823
22.2561 | -35.3385 | -32.8292
22.2622 | -35.3388 | -32.8354
22.2683 | -35.3392 | -32.8411
22.2744 | -35.3397 | -32.8472
22.2806 | -35.3401 | -32.8534
22.2867 | -35.3405 | -32.859
22.2928 | -35.3409 | -32.8655
22.2989 | -35.3414 | -32.8717
22.305 | -35.3418 | -32.878
22.3111 | -35.3422 | -32.8846
22.3172 | -35.3427 | -32.8913
22.3233 | -35.3432 | -32.8977
22.3294 | -35.3437 | -32.9041
22.3355 | -35.3441 | -32.9104
22.3416 | -35.3445 | -32.9162
22.3477 | -35.345 | -32.9225
22.3538 | -35.3454 | -32.9285
22.3599 | -35.3458 | -32.9345
22.366 | -35.3462 | -32.9406
22.3721 | -35.3466 | -32.9467
22.3782 | -35.347 | -32.9527
22.3843 | -35.3475 | -32.959
22.3904 | -35.348 | -32.9656
22.3965 | -35.3484 | -32.9721
22.4026 | -35.3488 | -32.9781
22.4087 | -35.3491 | -32.9837
22.4148 | -35.3495 | -32.9899
22.4209 | -35.35 | -32.9961
22.4271 | -35.3504 | -33.0029
22.4332 | -35.3508 | -33.0088
22.4393 | -35.3513 | -33.015
22.4454 | -35.3517 | -33.0207
22.4515 | -35.3522 | -33.0268
22.4576 | -35.3526 | -33.0331
22.4637 | -35.353 | -33.0397
22.4698 | -35.3535 | -33.0461
22.4759 | -35.3539 | -33.0515
22.482 | -35.3544 | -33.058
22.4881 | -35.3548 | -33.0645
22.4942 | -35.3553 | -33.071
22.5003 | -35.3557 | -33.0774
22.5064 | -35.356 | -33.0842
22.5125 | -35.3564 | -33.0909
22.5186 | -35.3568 | -33.0971
22.5247 | -35.3573 | -33.103
22.5308 | -35.3577 | -33.1092
22.5369 | -35.3582 | -33.1159
22.543 | -35.3586 | -33.1223
22.5491 | -35.359 | -33.1281
22.5552 | -35.3594 | -33.1347
22.5613 | -35.3599 | -33.1408
22.5675 | -35.3603 | -33.1471
22.5736 | -35.3606 | -33.1536
22.5797 | -35.361 | -33.1594
22.5858 | -35.3614 | -33.1653
22.5919 | -35.3618 | -33.1711
22.598 | -35.3622 | -33.1773
22.6041 | -35.3626 | -33.1837
22.6102 | -35.363 | -33.1901
22.6163 | -35.3634 | -33.1962
22.6224 | -35.3639 | -33.2021
22.6285 | -35.3644 | -33.2087
22.6346 | -35.3648 | -33.2146
22.6407 | -35.3652 | -33.2209
22.6468 | -35.3657 | -33.227
22.6529 | -35.366 | -33.2328
22.659 | -35.3664 | -33.2386
22.6651 | -35.3668 | -33.2449
22.6712 | -35.3672 | -33.2507
22.6773 | -35.3676 | -33.2562
22.6834 | -35.368 | -33.2624
22.6895 | -35.3684 | -33.2686
22.6956 | -35.3688 | -33.2744
22.7017 | -35.3692 | -33.2814
22.7079 | -35.3696 | -33.2879
22.714 | -35.3701 | -33.2939
22.7201 | -35.3705 | -33.3002
22.7262 | -35.3709 | -33.3063
22.7323 | -35.3713 | -33.3121
22.7384 | -35.3717 | -33.3188
22.7445 | -35.3722 | -33.3252
22.7506 | -35.3726 | -33.3312
22.7567 | -35.3731 | -33.3372
22.7628 | -35.3734 | -33.3432
22.7689 | -35.3738 | -33.3495
22.775 | -35.3743 | -33.3558
22.7811 | -35.3747 | -33.3617
22.7872 | -35.3751 | -33.3678
22.7933 | -35.3754 | -33.373
22.7994 | -35.3759 | -33.3789
22.8055 | -35.3762 | -33.3854
22.8116 | -35.3766 | -33.3918
22.8177 | -35.377 | -33.3979
22.8238 | -35.3774 | -33.4042
22.8299 | -35.3778 | -33.4107
22.836 | -35.3782 | -33.4175
22.8421 | -35.3786 | -33.4234
22.8482 | -35.379 | -33.4299
22.8544 | -35.3794 | -33.4365
22.8605 | -35.3799 | -33.4427
22.8666 | -35.3804 | -33.4493
22.8727 | -35.3809 | -33.4557
22.8788 | -35.3813 | -33.4617
22.8849 | -35.3817 | -33.4676
22.891 | -35.3821 | -33.4734
22.8971 | -35.3826 | -33.4797
22.9032 | -35.3829 | -33.4856
22.9093 | -35.3832 | -33.4914
22.9154 | -35.3837 | -33.498
22.9215 | -35.3841 | -33.5048
22.9276 | -35.3845 | -33.5111
22.9337 | -35.3849 | -33.5173
22.9398 | -35.3853 | -33.5238
22.9459 | -35.3857 | -33.5299
22.952 | -35.386 | -33.5362
22.9581 | -35.3864 | -33.5423
22.9642 | -35.3868 | -33.5485
22.9703 | -35.3872 | -33.555
22.9764 | -35.3876 | -33.5613
22.9825 | -35.388 | -33.5668
22.9886 | -35.3884 | -33.5731
22.9948 | -35.3888 | -33.5791
23.0009 | -35.3892 | -33.5852
23.007 | -35.3897 | -33.5919
23.0131 | -35.3901 | -33.5985
23.0192 | -35.3905 | -33.6051
23.0253 | -35.391 | -33.6109
23.0314 | -35.3914 | -33.6168
23.0375 | -35.3918 | -33.6234
23.0436 | -35.3922 | -33.6297
23.0497 | -35.3927 | -33.6357
23.0558 | -35.3931 | -33.6416
23.0619 | -35.3935 | -33.6479
23.068 | -35.3939 | -33.6537
23.0741 | -35.3942 | -33.6593
23.0802 | -35.3945 | -33.6652
23.0863 | -35.3948 | -33.6713
23.0924 | -35.3952 | -33.6774
23.0985 | -35.3956 | -33.6835
23.1046 | -35.396 | -33.6895
23.1107 | -35.3963 | -33.6963
23.1168 | -35.3966 | -33.7022
23.1229 | -35.397 | -33.7084
23.129 | -35.3974 | -33.7145
23.1351 | -35.3978 | -33.7204
23.1413 | -35.3982 | -33.7267
23.1474 | -35.3985 | -33.7324
23.1535 | -35.3989 | -33.7386
23.1596 | -35.3993 | -33.7447
23.1657 | -35.3996 | -33.7509
23.1718 | -35.4001 | -33.7576
23.1779 | -35.4005 | -33.7637
23.184 | -35.4007 | -33.7688
23.1901 | -35.4011 | -33.7751
23.1962 | -35.4015 | -33.781
23.2023 | -35.402 | -33.7873
23.2084 | -35.4024 | -33.7934
23.2145 | -35.4028 | -33.8
23.2206 | -35.4032 | -33.8064
23.2267 | -35.4036 | -33.8126
23.2328 | -35.4039 | -33.8186
23.2389 | -35.4043 | -33.8254
23.245 | -35.4047 | -33.8319
23.2511 | -35.4051 | -33.8371
23.2572 | -35.4056 | -33.8435
23.2633 | -35.406 | -33.8495
23.2694 | -35.4064 | -33.8557
23.2755 | -35.4068 | -33.862
23.2817 | -35.4072 | -33.8681
23.2878 | -35.4075 | -33.8744
23.2939 | -35.4079 | -33.8802
23.3 | -35.4083 | -33.8865
23.3061 | -35.4087 | -33.8923
23.3122 | -35.4091 | -33.8986
23.3183 | -35.4095 | -33.9051
23.3244 | -35.4098 | -33.9115
23.3305 | -35.4102 | -33.9175
23.3366 | -35.4105 | -33.9246
23.3427 | -35.411 | -33.9306
23.3488 | -35.4114 | -33.9365
23.3549 | -35.4117 | -33.9428
23.361 | -35.4121 | -33.9492
23.3671 | -35.4126 | -33.9553
23.3732 | -35.413 | -33.9612
23.3793 | -35.4134 | -33.9675
23.3854 | -35.4138 | -33.974
23.3915 | -35.4142 | -33.9795
23.3976 | -35.4146 | -33.9857
23.4037 | -35.415 | -33.9912
23.4098 | -35.4153 | -33.9971
23.4159 | -35.4158 | -34.0035
23.422 | -35.4161 | -34.0097
23.4282 | -35.4164 | -34.0155
23.4343 | -35.4168 | -34.0217
23.4404 | -35.4173 | -34.028
23.4465 | -35.4176 | -34.0338
23.4526 | -35.418 | -34.0405
23.4587 | -35.4182 | -34.0463
23.4648 | -35.4185 | -34.0529
23.4709 | -35.4188 | -34.0588
23.477 | -35.4191 | -34.0651
23.4831 | -35.4196 | -34.0714
23.4892 | -35.42 | -34.0774
23.4953 | -35.4204 | -34.0837
23.5014 | -35.4207 | -34.0908
23.5075 | -35.4211 | -34.0968
23.5136 | -35.4214 | -34.1035
23.5197 | -35.4219 | -34.1094
23.5258 | -35.4223 | -34.1154
23.5319 | -35.4227 | -34.1217
23.538 | -35.4231 | -34.1281
23.5441 | -35.4235 | -34.1342
23.5502 | -35.4238 | -34.1406
23.5563 | -35.4242 | -34.1466
23.5624 | -35.4246 | -34.1525
23.5686 | -35.4249 | -34.1582
23.5747 | -35.4252 | -34.1636
23.5808 | -35.4256 | -34.17
23.5869 | -35.426 | -34.1762
23.593 | -35.4263 | -34.1825
23.5991 | -35.4267 | -34.1888
23.6052 | -35.427 | -34.1953
23.6113 | -35.4274 | -34.2012
23.6174 | -35.4278 | -34.2071
23.6235 | -35.4281 | -34.2132
23.6296 | -35.4285 | -34.2194
23.6357 | -35.4288 | -34.2261
23.6418 | -35.4292 | -34.2317
23.6479 | -35.4296 | -34.2377
23.654 | -35.43 | -34.2439
23.6601 | -35.4304 | -34.25
23.6662 | -35.4307 | -34.2565
23.6723 | -35.4311 | -34.2623
23.6784 | -35.4315 | -34.2688
23.6845 | -35.4319 | -34.2753
23.6906 | -35.4323 | -34.2814
23.6967 | -35.4326 | -34.2871
23.7028 | -35.433 | -34.2934
23.7089 | -35.4334 | -34.2998
23.7151 | -35.4338 | -34.306
23.7212 | -35.4342 | -34.3127
23.7273 | -35.4345 | -34.3188
23.7334 | -35.4349 | -34.3255
23.7395 | -35.4351 | -34.3313
23.7456 | -35.4354 | -34.3375
23.7517 | -35.4357 | -34.3436
23.7578 | -35.4361 | -34.3499
23.7639 | -35.4365 | -34.3563
23.77 | -35.4369 | -34.3617
23.7761 | -35.4372 | -34.368
23.7822 | -35.4377 | -34.3742
23.7883 | -35.438 | -34.3805
23.7944 | -35.4384 | -34.387
23.8005 | -35.4388 | -34.3933
23.8066 | -35.4392 | -34.4
23.8127 | -35.4396 | -34.4059
23.8188 | -35.4399 | -34.4121
23.8249 | -35.4402 | -34.4181
23.831 | -35.4406 | -34.4243
23.8371 | -35.441 | -34.4303
23.8432 | -35.4414 | -34.4365
23.8493 | -35.4417 | -34.4425
23.8555 | -35.4421 | -34.4489
23.8616 | -35.4425 | -34.4551
23.8677 | -35.4429 | -34.4615
23.8738 | -35.4432 | -34.4679
23.8799 | -35.4435 | -34.4743
23.886 | -35.4438 | -34.4808
23.8921 | -35.4441 | -34.4875
23.8982 | -35.4445 | -34.4935
23.9043 | -35.4449 | -34.5001
23.9104 | -35.4452 | -34.5059
23.9165 | -35.4455 | -34.512
23.9226 | -35.446 | -34.5182
23.9287 | -35.4463 | -34.524
23.9348 | -35.4466 | -34.5302
23.9409 | -35.447 | -34.5362
23.947 | -35.4473 | -34.5423
23.9531 | -35.4477 | -34.5486
23.9592 | -35.4481 | -34.5548
23.9653 | -35.4485 | -34.5608
23.9714 | -35.4489 | -34.5667
23.9775 | -35.4493 | -34.5729
23.9836 | -35.4496 | -34.5793
23.9897 | -35.45 | -34.5852
23.9958 | -35.4504 | -34.5915
24.002 | -35.4508 | -34.5975
24.0081 | -35.451 | -34.6035
24.0142 | -35.4514 | -34.6098
24.0203 | -35.4517 | -34.6158
24.0264 | -35.4519 | -34.6213
24.0325 | -35.4523 | -34.6271
24.0386 | -35.4527 | -34.6338
24.0447 | -35.4531 | -34.6396
24.0508 | -35.4534 | -34.6459
24.0569 | -35.4538 | -34.6522
24.063 | -35.4541 | -34.6583
24.0691 | -35.4545 | -34.6643
24.0752 | -35.4549 | -34.6705
24.0813 | -35.4553 | -34.6765
24.0874 | -35.4556 | -34.6825
24.0935 | -35.456 | -34.6887
24.0996 | -35.4564 | -34.6947
24.1057 | -35.4567 | -34.7012
24.1118 | -35.4571 | -34.7075
24.1179 | -35.4574 | -34.7141
24.124 | -35.4577 | -34.7196
24.1301 | -35.4581 | -34.7255
24.1362 | -35.4585 | -34.7316
24.1424 | -35.4589 | -34.7377
24.1485 | -35.4592 | -34.7438
24.1546 | -35.4595 | -34.7503
24.1607 | -35.4598 | -34.7565
24.1668 | -35.4601 | -34.7626
24.1729 | -35.4605 | -34.7684
24.179 | -35.4609 | -34.7744
24.1851 | -35.4612 | -34.7808
24.1912 | -35.4616 | -34.7868
24.1973 | -35.462 | -34.7931
24.2034 | -35.4623 | -34.799
24.2095 | -35.4627 | -34.8051
24.2156 | -35.4631 | -34.8114
24.2217 | -35.4635 | -34.818
24.2278 | -35.4639 | -34.8241
24.2339 | -35.4643 | -34.8299
24.24 | -35.4646 | -34.8356
24.2461 | -35.4649 | -34.8419
24.2522 | -35.4653 | -34.8484
24.2583 | -35.4657 | -34.8551
24.2644 | -35.4661 | -34.8619
24.2705 | -35.4664 | -34.8682
24.2766 | -35.4667 | -34.8746
24.2827 | -35.4671 | -34.8808
24.2889 | -35.4675 | -34.8868
24.295 | -35.4678 | -34.8931
24.3011 | -35.4682 | -34.8992
24.3072 | -35.4686 | -34.9048
24.3133 | -35.469 | -34.9107
24.3194 | -35.4693 | -34.9171
24.3255 | -35.4696 | -34.9235
24.3316 | -35.47 | -34.9296
24.3377 | -35.4704 | -34.9361
24.3438 | -35.4706 | -34.9425
24.3499 | -35.4709 | -34.9485
24.356 | -35.4712 | -34.9549
24.3621 | -35.4716 | -34.9613
24.3682 | -35.4719 | -34.967
24.3743 | -35.4723 | -34.9733
24.3804 | -35.4726 | -34.9793
24.3865 | -35.4729 | -34.9856
24.3926 | -35.4733 | -34.9918
24.3987 | -35.4736 | -34.9977
24.4048 | -35.474 | -35.0037
24.4109 | -35.4743 | -35.0096
24.417 | -35.4747 | -35.0154
24.4231 | -35.475 | -35.022
24.4293 | -35.4754 | -35.0282
24.4354 | -35.4758 | -35.0342
24.4415 | -35.4762 | -35.0401
24.4476 | -35.4766 | -35.0461
24.4537 | -35.477 | -35.052
24.4598 | -35.4774 | -35.0585
24.4659 | -35.4778 | -35.0646
24.472 | -35.4782 | -35.0709
24.4781 | -35.4785 | -35.0772
24.4842 | -35.4788 | -35.0825
24.4903 | -35.4791 | -35.0881
24.4964 | -35.4794 | -35.0945
24.5025 | -35.4798 | -35.1004
24.5086 | -35.4802 | -35.1066
24.5147 | -35.4806 | -35.112
24.5208 | -35.481 | -35.1175
24.5269 | -35.4813 | -35.1229
24.533 | -35.4817 | -35.1294
24.5391 | -35.4821 | -35.1357
24.5452 | -35.4824 | -35.1423
24.5513 | -35.4827 | -35.1484
24.5574 | -35.4831 | -35.1545
24.5635 | -35.4835 | -35.1607
24.5696 | -35.4838 | -35.1672
24.5758 | -35.4842 | -35.1737
24.5819 | -35.4846 | -35.1797
24.588 | -35.485 | -35.1852
24.5941 | -35.4854 | -35.1912
24.6002 | -35.4858 | -35.1975
24.6063 | -35.4861 | -35.2045
24.6124 | -35.4864 | -35.2113
24.6185 | -35.4868 | -35.2176
24.6246 | -35.4871 | -35.2243
24.6307 | -35.4874 | -35.2306
24.6368 | -35.4878 | -35.2362
24.6429 | -35.4881 | -35.2424
24.649 | -35.4884 | -35.2485
24.6551 | -35.4888 | -35.2543
24.6612 | -35.4891 | -35.2607
24.6673 | -35.4895 | -35.2669
24.6734 | -35.4899 | -35.2732
24.6795 | -35.4902 | -35.2791
24.6856 | -35.4906 | -35.2852
24.6917 | -35.4909 | -35.2916
24.6978 | -35.4913 | -35.2973
24.7039 | -35.4916 | -35.3031
24.71 | -35.4919 | -35.3092
24.7162 | -35.4922 | -35.3154
24.7223 | -35.4925 | -35.3214
24.7284 | -35.4929 | -35.3276
24.7345 | -35.4933 | -35.3335
24.7406 | -35.4937 | -35.3397
24.7467 | -35.4941 | -35.3463
24.7528 | -35.4943 | -35.3524
24.7589 | -35.4947 | -35.3583
24.765 | -35.4951 | -35.3642
24.7711 | -35.4954 | -35.3703
24.7772 | -35.4958 | -35.3763
24.7833 | -35.4962 | -35.3825
24.7894 | -35.4965 | -35.389
24.7955 | -35.4969 | -35.3949
24.8016 | -35.4973 | -35.4012
24.8077 | -35.4977 | -35.4077
24.8138 | -35.498 | -35.4137
24.8199 | -35.4984 | -35.4201
24.826 | -35.4988 | -35.4257
24.8321 | -35.4991 | -35.4314
24.8382 | -35.4995 | -35.4375
24.8443 | -35.4998 | -35.4432
24.8504 | -35.5002 | -35.4494
24.8565 | -35.5006 | -35.4549
24.8627 | -35.501 | -35.4611
24.8688 | -35.5015 | -35.4669
24.8749 | -35.5018 | -35.4721
24.881 | -35.5022 | -35.4781
24.8871 | -35.5025 | -35.4842
24.8932 | -35.5029 | -35.4901
24.8993 | -35.5032 | -35.4961
24.9054 | -35.5036 | -35.5018
24.9115 | -35.504 | -35.5079
24.9176 | -35.5043 | -35.5138
24.9237 | -35.5047 | -35.52
24.9298 | -35.505 | -35.5261
24.9359 | -35.5053 | -35.532
24.942 | -35.5056 | -35.5376
24.9481 | -35.506 | -35.5435
24.9542 | -35.5063 | -35.55
24.9603 | -35.5067 | -35.5559
24.9664 | -35.507 | -35.5617
24.9725 | -35.5074 | -35.5679
24.9786 | -35.5077 | -35.5738
24.9847 | -35.508 | -35.5797
24.9908 | -35.5084 | -35.5859
24.9969 | -35.5087 | -35.5924
25.0031 | -35.509 | -35.5986
25.0092 | -35.5094 | -35.6039
25.0153 | -35.5097 | -35.6093
25.0214 | -35.51 | -35.6159
25.0275 | -35.5103 | -35.6216
25.0336 | -35.5107 | -35.6273
25.0397 | -35.5111 | -35.6337
25.0458 | -35.5114 | -35.6404
25.0519 | -35.5117 | -35.6471
25.058 | -35.5121 | -35.6536
25.0641 | -35.5125 | -35.6589
25.0702 | -35.5128 | -35.6647
25.0763 | -35.5131 | -35.6711
25.0824 | -35.5136 | -35.6772
25.0885 | -35.514 | -35.6829
25.0946 | -35.5144 | -35.6884
25.1007 | -35.5147 | -35.6942
25.1068 | -35.515 | -35.6998
25.1129 | -35.5154 | -35.7061
25.119 | -35.5158 | -35.7119
25.1251 | -35.5162 | -35.718
25.1312 | -35.5165 | -35.7237
25.1373 | -35.5168 | -35.7302
25.1435 | -35.5173 | -35.7368
25.1496 | -35.5177 | -35.743
25.1557 | -35.5181 | -35.7486
25.1618 | -35.5184 | -35.7551
25.1679 | -35.5187 | -35.7613
25.174 | -35.5191 | -35.767
25.1801 | -35.5195 | -35.7729
25.1862 | -35.5199 | -35.7792
25.1923 | -35.5202 | -35.7853
25.1984 | -35.5205 | -35.7918
25.2045 | -35.521 | -35.7979
25.2106 | -35.5213 | -35.8037
25.2167 | -35.5216 | -35.8102
25.2228 | -35.5219 | -35.8161
25.2289 | -35.5223 | -35.8221
25.235 | -35.5227 | -35.8278
25.2411 | -35.523 | -35.833
25.2472 | -35.5233 | -35.8382
25.2533 | -35.5236 | -35.8443
25.2594 | -35.524 | -35.8499
25.2655 | -35.5244 | -35.856
25.2716 | -35.5248 | -35.8617
25.2777 | -35.5251 | -35.8681
25.2838 | -35.5255 | -35.874
25.29 | -35.5258 | -35.88
25.2961 | -35.5263 | -35.8862
25.3022 | -35.5266 | -35.8919
25.3083 | -35.5269 | -35.8975
25.3144 | -35.5273 | -35.9042
25.3205 | -35.5276 | -35.9096
25.3266 | -35.528 | -35.9158
25.3327 | -35.5284 | -35.9223
25.3388 | -35.5287 | -35.928
25.3449 | -35.529 | -35.934
25.351 | -35.5294 | -35.94
25.3571 | -35.5297 | -35.9458
25.3632 | -35.5302 | -35.9515
25.3693 | -35.5305 | -35.9578
25.3754 | -35.5309 | -35.9634
25.3815 | -35.5314 | -35.9692
25.3876 | -35.5318 | -35.9751
25.3937 | -35.5321 | -35.9811
25.3998 | -35.5325 | -35.9864
25.4059 | -35.5329 | -35.9929
25.412 | -35.5332 | -35.9989
25.4181 | -35.5336 | -36.0053
25.4242 | -35.5339 | -36.011
25.4304 | -35.5342 | -36.017
25.4365 | -35.5347 | -36.0225
25.4426 | -35.535 | -36.028
25.4487 | -35.5353 | -36.034
25.4548 | -35.5357 | -36.0395
25.4609 | -35.536 | -36.0454
25.467 | -35.5365 | -36.052
25.4731 | -35.5368 | -36.0577
25.4792 | -35.5371 | -36.0639
25.4853 | -35.5375 | -36.0695
25.4914 | -35.5378 | -36.0757
25.4975 | -35.5382 | -36.0814
25.5036 | -35.5387 | -36.0871
25.5097 | -35.5391 | -36.0931
25.5158 | -35.5394 | -36.0986
25.5219 | -35.5398 | -36.1043
25.528 | -35.5403 | -36.1101
25.5341 | -35.5406 | -36.116
25.5402 | -35.5411 | -36.1214
25.5463 | -35.5415 | -36.1275
25.5524 | -35.5418 | -36.1339
25.5585 | -35.5422 | -36.1395
25.5646 | -35.5425 | -36.1456
25.5707 | -35.5429 | -36.1509
25.5769 | -35.5434 | -36.1569
25.583 | -35.5438 | -36.1631
25.5891 | -35.5442 | -36.1696
25.5952 | -35.5445 | -36.1755
25.6013 | -35.5449 | -36.1813
25.6074 | -35.5454 | -36.1872
25.6135 | -35.5458 | -36.1927
25.6196 | -35.5462 | -36.1984
25.6257 | -35.5465 | -36.204
25.6318 | -35.5469 | -36.2099
25.6379 | -35.5472 | -36.216
25.644 | -35.5477 | -36.2226
25.6501 | -35.5481 | -36.2284
25.6562 | -35.5483 | -36.2337
25.6623 | -35.5488 | -36.2394
25.6684 | -35.5492 | -36.2454
25.6745 | -35.5496 | -36.2506
25.6806 | -35.55 | -36.2559
25.6867 | -35.5504 | -36.2624
25.6928 | -35.5508 | -36.268
25.6989 | -35.5512 | -36.2739
25.705 | -35.5515 | -36.2794
25.7111 | -35.5519 | -36.2857
25.7173 | -35.5522 | -36.2914
25.7234 | -35.5525 | -36.2971
25.7295 | -35.5529 | -36.3028
25.7356 | -35.5532 | -36.309
25.7417 | -35.5536 | -36.3145
25.7478 | -35.554 | -36.3196
25.7539 | -35.5544 | -36.3251
25.76 | -35.5548 | -36.3315
25.7661 | -35.5551 | -36.3376
25.7722 | -35.5555 | -36.3434
25.7783 | -35.5559 | -36.3492
25.7844 | -35.5562 | -36.355
25.7905 | -35.5565 | -36.361
25.7966 | -35.5569 | -36.3668
25.8027 | -35.5572 | -36.3727
25.8088 | -35.5575 | -36.379
25.8149 | -35.5579 | -36.3846
25.821 | -35.5583 | -36.3903
25.8271 | -35.5587 | -36.3964
25.8332 | -35.5591 | -36.4024
25.8393 | -35.5595 | -36.4083
25.8454 | -35.5598 | -36.4138
25.8515 | -35.5603 | -36.4195
25.8576 | -35.5606 | -36.4252
25.8638 | -35.561 | -36.4316
25.8699 | -35.5614 | -36.4371
25.876 | -35.5618 | -36.4431
25.8821 | -35.5622 | -36.4489
25.8882 | -35.5626 | -36.4542
25.8943 | -35.5629 | -36.4604
25.9004 | -35.5633 | -36.4663
25.9065 | -35.5637 | -36.472
25.9126 | -35.564 | -36.4776
25.9187 | -35.5644 | -36.4828
25.9248 | -35.5648 | -36.4887
25.9309 | -35.5652 | -36.4943
25.937 | -35.5656 | -36.5002
25.9431 | -35.5659 | -36.506
25.9492 | -35.5664 | -36.5118
25.9553 | -35.5667 | -36.5177
25.9614 | -35.5672 | -36.5231
25.9675 | -35.5676 | -36.5287
25.9736 | -35.5679 | -36.5347
25.9797 | -35.5684 | -36.5406
25.9858 | -35.5687 | -36.5461
25.9919 | -35.5691 | -36.5521
25.998 | -35.5695 | -36.5575
26.0042 | -35.5699 | -36.5635
26.0103 | -35.5703 | -36.5696
26.0164 | -35.5706 | -36.5752
26.0225 | -35.571 | -36.5811
26.0286 | -35.5714 | -36.5869
26.0347 | -35.5718 | -36.5923
26.0408 | -35.5722 | -36.5984
26.0469 | -35.5725 | -36.6035
26.053 | -35.5729 | -36.6094
26.0591 | -35.5733 | -36.6151
26.0652 | -35.5737 | -36.6205
26.0713 | -35.5741 | -36.6262
26.0774 | -35.5745 | -36.6319
26.0835 | -35.5749 | -36.6376
26.0896 | -35.5752 | -36.6433
26.0957 | -35.5756 | -36.6488
26.1018 | -35.576 | -36.6541
26.1079 | -35.5764 | -36.6598
26.114 | -35.5768 | -36.666
26.1201 | -35.5771 | -36.6718
26.1262 | -35.5774 | -36.6777
26.1323 | -35.5778 | -36.6831
26.1384 | -35.5782 | -36.6881
26.1445 | -35.5786 | -36.694
26.1507 | -35.579 | -36.6993
26.1568 | -35.5794 | -36.7052
26.1629 | -35.5798 | -36.7106
26.169 | -35.5801 | -36.716
26.1751 | -35.5805 | -36.7209
26.1812 | -35.5808 | -36.7264
26.1873 | -35.5812 | -36.7322
26.1934 | -35.5816 | -36.7378
26.1995 | -35.582 | -36.743
26.2056 | -35.5824 | -36.7479
26.2117 | -35.5828 | -36.7537
26.2178 | -35.5831 | -36.7592
26.2239 | -35.5835 | -36.7644
26.23 | -35.5839 | -36.7703
26.2361 | -35.5842 | -36.776
26.2422 | -35.5846 | -36.782
26.2483 | -35.585 | -36.7876
26.2544 | -35.5854 | -36.7927
26.2605 | -35.5858 | -36.7984
26.2666 | -35.5862 | -36.8035
26.2727 | -35.5866 | -36.8097
26.2788 | -35.5871 | -36.8151
26.2849 | -35.5874 | -36.8209
26.2911 | -35.5877 | -36.8264
26.2972 | -35.5881 | -36.8319
26.3033 | -35.5885 | -36.8376
26.3094 | -35.5889 | -36.8433
26.3155 | -35.5894 | -36.8487
26.3216 | -35.5897 | -36.8544
26.3277 | -35.5901 | -36.8599
26.3338 | -35.5905 | -36.8655
26.3399 | -35.5909 | -36.8712
26.346 | -35.5912 | -36.8764
26.3521 | -35.5916 | -36.8825
26.3582 | -35.592 | -36.8882
26.3643 | -35.5923 | -36.8941
26.3704 | -35.5927 | -36.9001
26.3765 | -35.5931 | -36.9057
26.3826 | -35.5936 | -36.9106
26.3887 | -35.5939 | -36.916
26.3948 | -35.5943 | -36.9217
26.4009 | -35.5947 | -36.9272
26.407 | -35.595 | -36.9325
26.4131 | -35.5954 | -36.9378
26.4192 | -35.5956 | -36.9435
26.4253 | -35.596 | -36.949
26.4314 | -35.5964 | -36.9542
26.4376 | -35.5968 | -36.96
26.4437 | -35.5971 | -36.9655
26.4498 | -35.5975 | -36.9712
26.4559 | -35.5979 | -36.9768
26.462 | -35.5983 | -36.9824
26.4681 | -35.5987 | -36.9878
26.4742 | -35.5991 | -36.9933
26.4803 | -35.5995 | -36.9987
26.4864 | -35.5999 | -37.0041
26.4925 | -35.6003 | -37.0095
26.4986 | -35.6007 | -37.0153
26.5047 | -35.601 | -37.0209
26.5108 | -35.6014 | -37.0262
26.5169 | -35.6018 | -37.0319
26.523 | -35.6022 | -37.0381
26.5291 | -35.6026 | -37.0437
26.5352 | -35.6029 | -37.0492
26.5413 | -35.6034 | -37.0546
26.5474 | -35.6038 | -37.0602
26.5535 | -35.6042 | -37.0658
26.5596 | -35.6045 | -37.0709
26.5657 | -35.6049 | -37.0766
26.5718 | -35.6053 | -37.0823
26.578 | -35.6057 | -37.088
26.5841 | -35.6061 | -37.0935
26.5902 | -35.6065 | -37.099
26.5963 | -35.6068 | -37.1043
26.6024 | -35.6072 | -37.1097
26.6085 | -35.6076 | -37.1156
26.6146 | -35.608 | -37.1212
26.6207 | -35.6084 | -37.1268
26.6268 | -35.6087 | -37.1322
26.6329 | -35.6092 | -37.1377
26.639 | -35.6095 | -37.1433
26.6451 | -35.6099 | -37.1488
26.6512 | -35.6103 | -37.154
26.6573 | -35.6107 | -37.1596
26.6634 | -35.6111 | -37.1644
26.6695 | -35.6114 | -37.1699
26.6756 | -35.6118 | -37.1753
26.6817 | -35.6121 | -37.181
26.6878 | -35.6125 | -37.1867
26.6939 | -35.6129 | -37.1922
26.7 | -35.6133 | -37.1973
26.7061 | -35.6137 | -37.2028
26.7122 | -35.614 | -37.2084
26.7183 | -35.6144 | -37.2138
26.7245 | -35.6148 | -37.2191
26.7306 | -35.6152 | -37.2246
26.7367 | -35.6156 | -37.2304
26.7428 | -35.6161 | -37.236
26.7489 | -35.6164 | -37.2413
26.755 | -35.6168 | -37.2469
26.7611 | -35.6171 | -37.2527
26.7672 | -35.6175 | -37.2581
26.7733 | -35.6179 | -37.2636
26.7794 | -35.6183 | -37.2693
26.7855 | -35.6188 | -37.2749
26.7916 | -35.6191 | -37.2802
26.7977 | -35.6195 | -37.2859
26.8038 | -35.6199 | -37.2911
26.8099 | -35.6202 | -37.2965
26.816 | -35.6206 | -37.3024
26.8221 | -35.621 | -37.3084
26.8282 | -35.6214 | -37.3134
26.8343 | -35.6217 | -37.3193
26.8404 | -35.6221 | -37.3246
26.8465 | -35.6225 | -37.3306
26.8526 | -35.6228 | -37.3363
26.8587 | -35.6232 | -37.3414
26.8649 | -35.6236 | -37.3465
26.871 | -35.624 | -37.3524
26.8771 | -35.6243 | -37.3574
26.8832 | -35.6247 | -37.3628
26.8893 | -35.625 | -37.3685
26.8954 | -35.6254 | -37.3738
26.9015 | -35.6257 | -37.3794
26.9076 | -35.6261 | -37.3853
26.9137 | -35.6265 | -37.3905
26.9198 | -35.6269 | -37.3961
26.9259 | -35.6273 | -37.4009
26.932 | -35.6277 | -37.4062
26.9381 | -35.6281 | -37.4119
26.9442 | -35.6285 | -37.417
26.9503 | -35.6288 | -37.4224
26.9564 | -35.6293 | -37.4279
26.9625 | -35.6297 | -37.4334
26.9686 | -35.6301 | -37.4388
26.9747 | -35.6305 | -37.4442
26.9808 | -35.6309 | -37.4492
26.9869 | -35.6313 | -37.4545
26.993 | -35.6316 | -37.4597
26.9991 | -35.6321 | -37.465
27.0052 | -35.6325 | -37.4708
27.0114 | -35.6329 | -37.4766
27.0175 | -35.6333 | -37.4816
27.0236 | -35.6337 | -37.4871
27.0297 | -35.634 | -37.4927
27.0358 | -35.6343 | -37.4979
27.0419 | -35.6347 | -37.5032
27.048 | -35.6351 | -37.5089
27.0541 | -35.6354 | -37.5141
27.0602 | -35.6358 | -37.5196
27.0663 | -35.6362 | -37.525
27.0724 | -35.6367 | -37.5302
27.0785 | -35.6371 | -37.5358
27.0846 | -35.6375 | -37.5412
27.0907 | -35.6379 | -37.5467
27.0968 | -35.6383 | -37.5526
27.1029 | -35.6388 | -37.558
27.109 | -35.6391 | -37.5632
27.1151 | -35.6396 | -37.5692
27.1212 | -35.6399 | -37.5741
27.1273 | -35.6403 | -37.5795
27.1334 | -35.6407 | -37.5849
27.1395 | -35.6411 | -37.5903
27.1456 | -35.6415 | -37.5958
27.1518 | -35.6419 | -37.6011
27.1579 | -35.6423 | -37.6061
27.164 | -35.6427 | -37.6118
27.1701 | -35.643 | -37.617
27.1762 | -35.6434 | -37.6222
27.1823 | -35.6437 | -37.6277
27.1884 | -35.644 | -37.6334
27.1945 | -35.6444 | -37.639
27.2006 | -35.6448 | -37.6441
27.2067 | -35.6451 | -37.6493
27.2128 | -35.6455 | -37.6547
27.2189 | -35.6459 | -37.66
27.225 | -35.6463 | -37.6655
27.2311 | -35.6467 | -37.6711
27.2372 | -35.6471 | -37.6769
27.2433 | -35.6474 | -37.6819
27.2494 | -35.6478 | -37.6868
27.2555 | -35.6482 | -37.6923
27.2616 | -35.6486 | -37.6976
27.2677 | -35.6489 | -37.7024
27.2738 | -35.6494 | -37.7073
27.2799 | -35.6498 | -37.7128
27.286 | -35.6502 | -37.718
27.2921 | -35.6505 | -37.7236
27.2983 | -35.6509 | -37.729
27.3044 | -35.6513 | -37.7348
27.3105 | -35.6517 | -37.7409
27.3166 | -35.6521 | -37.7463
27.3227 | -35.6525 | -37.7511
27.3288 | -35.6529 | -37.7564
27.3349 | -35.6532 | -37.7616
27.341 | -35.6536 | -37.7668
27.3471 | -35.6539 | -37.7721
27.3532 | -35.6543 | -37.7774
27.3593 | -35.6548 | -37.783
27.3654 | -35.6553 | -37.7883
27.3715 | -35.6557 | -37.7934
27.3776 | -35.6561 | -37.799
27.3837 | -35.6565 | -37.8044
27.3898 | -35.6569 | -37.8098
27.3959 | -35.6573 | -37.8153
27.402 | -35.6578 | -37.8201
27.4081 | -35.6582 | -37.8252
27.4142 | -35.6586 | -37.8302
27.4203 | -35.659 | -37.8353
27.4264 | -35.6593 | -37.841
27.4325 | -35.6597 | -37.8461
27.4387 | -35.6601 | -37.8509
27.4448 | -35.6605 | -37.8556
27.4509 | -35.6609 | -37.8613
27.457 | -35.6613 | -37.866
27.4631 | -35.6617 | -37.8714
27.4692 | -35.6621 | -37.8766
27.4753 | -35.6625 | -37.8817
27.4814 | -35.6629 | -37.8866
27.4875 | -35.6633 | -37.8924
27.4936 | -35.6637 | -37.8975
27.4997 | -35.664 | -37.9027
27.5058 | -35.6644 | -37.9081
27.5119 | -35.6647 | -37.9139
27.518 | -35.6651 | -37.9188
27.5241 | -35.6655 | -37.9236
27.5302 | -35.6659 | -37.9286
27.5363 | -35.6663 | -37.9343
27.5424 | -35.6667 | -37.9395
27.5485 | -35.6671 | -37.9445
27.5546 | -35.6675 | -37.95
27.5607 | -35.6679 | -37.9552
27.5668 | -35.6683 | -37.9605
27.5729 | -35.6687 | -37.9658
27.5791 | -35.6691 | -37.9705
27.5852 | -35.6696 | -37.9758
27.5913 | -35.67 | -37.9811
27.5974 | -35.6704 | -37.9866
27.6035 | -35.6708 | -37.9918
27.6096 | -35.6711 | -37.9968
27.6157 | -35.6715 | -38.0018
27.6218 | -35.6719 | -38.0071
27.6279 | -35.6722 | -38.0129
27.634 | -35.6725 | -38.0184
27.6401 | -35.6729 | -38.024
27.6462 | -35.6733 | -38.0295
27.6523 | -35.6737 | -38.0339
27.6584 | -35.6741 | -38.0391
27.6645 | -35.6745 | -38.0442
27.6706 | -35.6749 | -38.0492
27.6767 | -35.6754 | -38.054
27.6828 | -35.6759 | -38.0588
27.6889 | -35.6763 | -38.0642
27.695 | -35.6768 | -38.0693
27.7011 | -35.6772 | -38.0747
27.7072 | -35.6776 | -38.0802
27.7133 | -35.6779 | -38.0857
27.7194 | -35.6784 | -38.0902
27.7256 | -35.6787 | -38.0953
27.7317 | -35.6791 | -38.1003
27.7378 | -35.6794 | -38.1055
27.7439 | -35.6798 | -38.1107
27.75 | -35.6802 | -38.1162
27.7561 | -35.6806 | -38.1213
27.7622 | -35.681 | -38.1262
27.7683 | -35.6814 | -38.1316
27.7744 | -35.6818 | -38.1365
27.7805 | -35.6822 | -38.1417
27.7866 | -35.6825 | -38.1465
27.7927 | -35.6828 | -38.1511
27.7988 | -35.6832 | -38.1565
27.8049 | -35.6836 | -38.1618
27.811 | -35.6839 | -38.1673
27.8171 | -35.6843 | -38.1724
27.8232 | -35.6847 | -38.1775
27.8293 | -35.685 | -38.1833
27.8354 | -35.6854 | -38.1886
27.8415 | -35.6858 | -38.1933
27.8476 | -35.6861 | -38.1986
27.8537 | -35.6864 | -38.2041
27.8598 | -35.6867 | -38.2089
27.866 | -35.6871 | -38.2141
27.8721 | -35.6875 | -38.2199
27.8782 | -35.6879 | -38.2256
27.8843 | -35.6883 | -38.2311
27.8904 | -35.6888 | -38.2358
27.8965 | -35.689 | -38.241
27.9026 | -35.6894 | -38.2456
27.9087 | -35.6897 | -38.2502
27.9148 | -35.6901 | -38.2547
27.9209 | -35.6904 | -38.2602
27.927 | -35.6909 | -38.2651
27.9331 | -35.6912 | -38.27
27.9392 | -35.6917 | -38.2746
27.9453 | -35.6921 | -38.2792
27.9514 | -35.6924 | -38.2846
27.9575 | -35.6927 | -38.2897
27.9636 | -35.6931 | -38.2951
27.9697 | -35.6935 | -38.2997
27.9758 | -35.694 | -38.3052
27.9819 | -35.6943 | -38.3105
27.988 | -35.6947 | -38.3155
27.9941 | -35.6951 | -38.3207
28.0002 | -35.6955 | -38.3254
28.0063 | -35.6958 | -38.3305
28.0125 | -35.6962 | -38.3356
28.0186 | -35.6966 | -38.3403
28.0247 | -35.697 | -38.3453
28.0308 | -35.6974 | -38.3505
28.0369 | -35.6979 | -38.355
28.043 | -35.6982 | -38.3604
28.0491 | -35.6986 | -38.3659
28.0552 | -35.699 | -38.3713
28.0613 | -35.6993 | -38.3765
28.0674 | -35.6997 | -38.3815
28.0735 | -35.6999 | -38.3863

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5116 | -43.0982
21.4908 | -35.5194 | -35.2138
16.4948 | -34.8993 | -27.7885
12.6602 | -34.5379 | -21.4834
9.7171 | -34.3229 | -16.5688
7.4582 | -34.1695 | -12.5774
5.7244 | -34.0628 | -9.3494
4.3936 | -33.9926 | -6.7364
3.3722 | -33.9777 | -4.6804
2.5883 | -33.9711 | -3.1247
1.9866 | -33.9692 | -1.8791
1.5248 | -33.9715 | -0.87949
1.1703 | -33.9819 | -0.23058
1 | -33.9763 | 0.094005
0.89822 | -33.9698 | 0.36737
0.68941 | -33.9509 | 0.94954
0.52919 | -33.9515 | 1.4024
0.40612 | -33.9578 | 2.0103
0.31171 | -33.9421 | 2.4677
0.23928 | -33.9563 | 3.1873
0.18358 | -33.9589 | 3.6048
0.14091 | -33.9821 | 4.4765
0.10819 | -34.0003 | 5.636
0.083017 | -34.0275 | 7.1052
0.063705 | -34.0728 | 8.9693
0.048923 | -34.14 | 11.4473
0.037479 | -34.2335 | 14.6324
0.028801 | -34.4059 | 18.6512
0.022078 | -34.7033 | 24.0704
0.017023 | -35.187 | 30.5363
0.013018 | -35.9977 | 38.2407
0.010014 | -37.1208 | 46.1684

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003711 | -10.9347
0.007423 | -35.3639
0.011134 | -38.9628
0.014846 | -42.7858
0.018557 | -44.4622
0.022269 | -46.1412
0.02598 | -48.9745
0.029691 | -51.6776
0.033403 | -52.2744
0.037114 | -52.0831
0.040826 | -53.6336
0.044537 | -54.4426
0.048249 | -55.705
0.05196 | -56.9252
0.055672 | -56.4482
0.059383 | -56.9321
0.063094 | -58.1919
0.066806 | -58.6899
0.070517 | -60.3175
0.074229 | -61.1354
0.07794 | -60.7128
0.081652 | -61.1147
0.085363 | -62.1911
0.089074 | -62.6012
0.092786 | -63.0022
0.096497 | -62.9281
0.10021 | -63.1652
0.10392 | -63.7457
0.10763 | -64.1324
0.11134 | -65.9262
0.11506 | -65.1413
0.11877 | -64.3328
0.12248 | -65.5308
0.12619 | -65.9437
0.1299 | -65.7902
0.13361 | -66.389
0.13732 | -67.4768
0.14103 | -68.2687
0.14475 | -67.6248
0.14846 | -68.5427
0.15217 | -68.4779
0.15588 | -68.2836
0.15959 | -68.6667
0.1633 | -68.9421
0.16701 | -69.2939
0.17073 | -69.5885
0.17444 | -69.3928
0.17815 | -70.0524
0.18186 | -70.4728
0.18557 | -70.8552
0.18928 | -70.939
0.19299 | -70.7709
0.19671 | -70.5437
0.20042 | -70.6959
0.20413 | -71.6112
0.20784 | -70.6895
0.21155 | -71.5578
0.21526 | -72.2562
0.21898 | -71.7837
0.22269 | -72.1069
0.2264 | -71.6525
0.23011 | -71.4146
0.23382 | -72.0906
0.23753 | -72.6842
0.24124 | -72.9436
0.24496 | -73.5874
0.24867 | -74.1145
0.25238 | -73.7016
0.25609 | -73.5243
0.2598 | -75.3056
0.26351 | -75.2066
0.26722 | -74.5426
0.27093 | -74.6143
0.27465 | -75.4548
0.27836 | -74.3915
0.28207 | -74.7194
0.28578 | -75.2458
0.28949 | -75.7992
0.2932 | -75.6444
0.29691 | -75.834
0.30063 | -74.9248
0.30434 | -75.8248
0.30805 | -76.3288
0.31176 | -76.1358
0.31547 | -76.438
0.31918 | -75.498
0.32289 | -75.3473
0.32661 | -75.8628
0.33032 | -75.9075
0.33403 | -75.8037
0.33774 | -75.6618
0.34145 | -76.2108
0.34516 | -76.7739
0.34887 | -77.0934
0.35259 | -77.3749
0.3563 | -78.6326
0.36001 | -78.8233
0.36372 | -78.4728
0.36743 | -77.5789
0.37114 | -77.4526
0.37485 | -78.3347
0.37857 | -78.1368
0.38228 | -79.3331
0.38599 | -78.7122
0.3897 | -79.5973
0.39341 | -79.4362
0.39712 | -78.3285
0.40083 | -78.8227
0.40455 | -78.8134
0.40826 | -78.2865
0.41197 | -77.781
0.41568 | -78.0974
0.41939 | -78.6244
0.4231 | -79.0258
0.42681 | -78.6164
0.43053 | -78.7853
0.43424 | -78.8556
0.43795 | -79.2852
0.44166 | -78.7133
0.44537 | -78.2526
0.44908 | -78.548
0.4528 | -79.8422
0.45651 | -80.5289
0.46022 | -80.5946
0.46393 | -80.1543
0.46764 | -79.5512
0.47135 | -80.4068
0.47506 | -81.4417
0.47878 | -81.4492
0.48249 | -80.5625
0.4862 | -80.0881
0.48991 | -80.7187
0.49362 | -81.1628
0.49733 | -81.1449
0.50104 | -80.9094
0.50475 | -79.1218
0.50847 | -79.8379
0.51218 | -81.0764
0.51589 | -81.3367
0.5196 | -81.6755
0.52331 | -82.6819
0.52702 | -81.8568
0.53073 | -81.1603
0.53445 | -81.2568
0.53816 | -81.6149
0.54187 | -82.0476
0.54558 | -81.7567
0.54929 | -81.3968
0.553 | -81.6524
0.55671 | -82.1643
0.56043 | -82.4481
0.56414 | -82.1176
0.56785 | -81.3216
0.57156 | -81.1495
0.57527 | -82.0103
0.57898 | -83.0344
0.58269 | -83.0581
0.58641 | -82.3145
0.59012 | -81.8273
0.59383 | -82.0306
0.59754 | -82.7478
0.60125 | -83.2709
0.60496 | -83.056
0.60867 | -82.8161
0.61239 | -82.8057
0.6161 | -83.1502
0.61981 | -83.566
0.62352 | -83.7533
0.62723 | -83.4092
0.63094 | -82.9968
0.63466 | -82.9173
0.63837 | -82.9239
0.64208 | -82.5962
0.64579 | -81.9276
0.6495 | -82.0274
0.65321 | -82.989
0.65692 | -84.1448
0.66064 | -84.203
0.66435 | -83.4797
0.66806 | -83.2247
0.67177 | -83.8582
0.67548 | -84.6934
0.67919 | -84.8871
0.6829 | -84.7258
0.68662 | -84.4054
0.69033 | -84.3966
0.69404 | -84.6598
0.69775 | -84.4437
0.70146 | -83.7797
0.70517 | -83.66
0.70888 | -84.0571
0.7126 | -84.237
0.71631 | -84.1374
0.72002 | -84.0338
0.72373 | -84.1269
0.72744 | -84.4336
0.73115 | -84.9686
0.73486 | -86.0747
0.73858 | -86.0681
0.74229 | -84.7141
0.746 | -83.8219
0.74971 | -83.9081
0.75342 | -84.5263
0.75713 | -84.9142
0.76084 | -85.1472
0.76456 | -85.5464
0.76827 | -85.3726
0.77198 | -85.3931
0.77569 | -85.7663
0.7794 | -85.824
0.78311 | -85.7334
0.78682 | -85.6899
0.79054 | -85.5973
0.79425 | -85.7332
0.79796 | -85.3799
0.80167 | -84.6986
0.80538 | -84.5969
0.80909 | -85.2527
0.8128 | -85.6571
0.81652 | -85.3123
0.82023 | -84.802
0.82394 | -84.7248
0.82765 | -85.1994
0.83136 | -86.0206
0.83507 | -86.4427
0.83879 | -86.0475
0.8425 | -85.4323
0.84621 | -85.0943
0.84992 | -85.0043
0.85363 | -85.3383
0.85734 | -85.6364
0.86105 | -85.8464
0.86477 | -86.2662
0.86848 | -86.5495
0.87219 | -86.5021
0.8759 | -86.3379
0.87961 | -86.2598
0.88332 | -86.3175
0.88703 | -86.5216
0.89075 | -86.9233
0.89446 | -86.7079
0.89817 | -85.8505
0.90188 | -85.631
0.90559 | -85.9957
0.9093 | -86.0959
0.91301 | -86.0698
0.91673 | -86.3016
0.92044 | -86.9487
0.92415 | -87.2442
0.92786 | -87.1411
0.93157 | -87.1005
0.93528 | -86.9965
0.93899 | -86.6845
0.94271 | -86.4759
0.94642 | -86.4627
0.95013 | -86.3982
0.95384 | -86.3619
0.95755 | -86.1342
0.96126 | -86.0371
0.96497 | -86.435
0.96869 | -87.2944
0.9724 | -87.7656
0.97611 | -87.56
0.97982 | -87.2659
0.98353 | -87.4422
0.98724 | -87.6916
0.99095 | -87.9784
0.99467 | -88.1169
0.99838 | -88.252
1.0021 | -88.3852
1.0058 | -88.1228
1.0095 | -87.7153
1.0132 | -87.5123
1.0169 | -87.6156
1.0206 | -87.4714
1.0244 | -86.9331
1.0281 | -86.5234
1.0318 | -86.368
1.0355 | -86.659
1.0392 | -87.4425
1.0429 | -88.351
1.0466 | -88.5666
1.0503 | -88.3201
1.054 | -87.9642
1.0578 | -87.3811
1.0615 | -87.1326
1.0652 | -87.2076
1.0689 | -87.3722
1.0726 | -87.5478
1.0763 | -87.655
1.08 | -87.6946
1.0837 | -87.8104
1.0875 | -88.032
1.0912 | -88.2481
1.0949 | -88.3838
1.0986 | -88.4764
1.1023 | -88.5929
1.106 | -88.6861
1.1097 | -88.6533
1.1134 | -88.54
1.1171 | -88.4865
1.1209 | -88.533
1.1246 | -88.6363
1.1283 | -88.6975
1.132 | -88.6827
1.1357 | -88.653
1.1394 | -88.4911
1.1431 | -88.1737
1.1468 | -87.9743
1.1505 | -87.8323
1.1543 | -87.5525
1.158 | -87.4219
1.1617 | -87.7171
1.1654 | -88.1742
1.1691 | -88.3343
1.1728 | -88.2947
1.1765 | -88.3716
1.1802 | -88.5643
1.1839 | -88.8001
1.1877 | -89.0286
1.1914 | -89.1389
1.1951 | -89.153
1.1988 | -89.1669
1.2025 | -89.2889
1.2062 | -89.576
1.2099 | -89.9345
1.2136 | -90.1418
1.2174 | -90.1142
1.2211 | -89.8965
1.2248 | -89.5945
1.2285 | -89.439
1.2322 | -89.5691
1.2359 | -89.7357
1.2396 | -89.5596
1.2433 | -89.1107
1.247 | -88.7888
1.2508 | -88.6876
1.2545 | -88.6804
1.2582 | -88.7709
1.2619 | -88.9749
1.2656 | -89.146
1.2693 | -89.2323
1.273 | -89.4227
1.2767 | -89.6297
1.2804 | -89.535
1.2842 | -89.2993
1.2879 | -89.2431
1.2916 | -89.3956
1.2953 | -89.6392
1.299 | -89.9078
1.3027 | -90.1423
1.3064 | -90.2731
1.3101 | -90.3629
1.3138 | -90.446
1.3176 | -90.3928
1.3213 | -90.2004
1.325 | -90.0532
1.3287 | -90.0735
1.3324 | -90.2108
1.3361 | -90.2585
1.3398 | -89.9921
1.3435 | -89.564
1.3473 | -89.3207
1.351 | -89.3044
1.3547 | -89.3424
1.3584 | -89.3566
1.3621 | -89.3737
1.3658 | -89.4856
1.3695 | -89.734
1.3732 | -89.9553
1.3769 | -90.1581
1.3807 | -90.4746
1.3844 | -90.7436
1.3881 | -90.7426
1.3918 | -90.5217
1.3955 | -90.3177
1.3992 | -90.2336
1.4029 | -90.2343
1.4066 | -90.3296
1.4103 | -90.3523
1.4141 | -90.103
1.4178 | -89.7279
1.4215 | -89.5777
1.4252 | -89.8266
1.4289 | -90.2921
1.4326 | -90.7029
1.4363 | -91.0449
1.44 | -91.3253
1.4437 | -91.3835
1.4475 | -91.2482
1.4512 | -91.1784
1.4549 | -91.1582
1.4586 | -91.0105
1.4623 | -90.7993
1.466 | -90.6905
1.4697 | -90.6612
1.4734 | -90.6292
1.4772 | -90.6624
1.4809 | -90.7807
1.4846 | -90.7882
1.4883 | -90.7401
1.492 | -90.8574
1.4957 | -91.0071
1.4994 | -90.9381
1.5031 | -90.7553
1.5068 | -90.6797
1.5106 | -90.7629
1.5143 | -90.9261
1.518 | -91.0564
1.5217 | -91.1147
1.5254 | -91.2395
1.5291 | -91.5168
1.5328 | -91.8124
1.5365 | -91.9321
1.5402 | -91.831
1.544 | -91.5652
1.5477 | -91.2621
1.5514 | -91.1581
1.5551 | -91.367
1.5588 | -91.6298
1.5625 | -91.6358
1.5662 | -91.4644
1.5699 | -91.3412
1.5736 | -91.4202
1.5774 | -91.7385
1.5811 | -92.075
1.5848 | -92.1894
1.5885 | -92.0985
1.5922 | -91.9016
1.5959 | -91.801
1.5996 | -91.9383
1.6033 | -91.9709
1.6071 | -91.8986
1.6108 | -91.6989
1.6145 | -91.411
1.6182 | -91.1045
1.6219 | -90.9132
1.6256 | -90.9396
1.6293 | -91.1999
1.633 | -91.5795
1.6367 | -91.8862
1.6405 | -91.965
1.6442 | -91.8603
1.6479 | -91.7399
1.6516 | -91.7256
1.6553 | -91.8033
1.659 | -91.8999
1.6627 | -91.9588
1.6664 | -91.9415
1.6701 | -91.7973
1.6739 | -91.5788
1.6776 | -91.4187
1.6813 | -91.4495
1.685 | -91.6905
1.6887 | -92.0369
1.6924 | -92.295
1.6961 | -92.2915
1.6998 | -92.0209
1.7035 | -91.6721
1.7073 | -91.4673
1.711 | -91.4349
1.7147 | -91.4172
1.7184 | -91.3318
1.7221 | -91.2676
1.7258 | -91.3465
1.7295 | -91.5615
1.7332 | -91.8189
1.737 | -92.0552
1.7407 | -92.2835
1.7444 | -92.5234
1.7481 | -92.7345
1.7518 | -92.8817
1.7555 | -92.9524
1.7592 | -92.928
1.7629 | -92.8292
1.7666 | -92.724
1.7704 | -92.6683
1.7741 | -92.6265
1.7778 | -92.5111
1.7815 | -92.2916
1.7852 | -92.0631
1.7889 | -91.9098
1.7926 | -91.8261
1.7963 | -91.8018
1.8 | -91.8696
1.8038 | -92.0499
1.8075 | -92.2886
1.8112 | -92.4741
1.8149 | -92.5785
1.8186 | -92.6559
1.8223 | -92.7795
1.826 | -92.9223
1.8297 | -92.9899
1.8334 | -92.9007
1.8372 | -92.6438
1.8409 | -92.3131
1.8446 | -92.0993
1.8483 | -92.1372
1.852 | -92.4334
1.8557 | -92.8599
1.8594 | -93.2566
1.8631 | -93.5105
1.8669 | -93.6036
1.8706 | -93.6192
1.8743 | -93.6332
1.878 | -93.6213
1.8817 | -93.5433
1.8854 | -93.4155
1.8891 | -93.3471
1.8928 | -93.4176
1.8965 | -93.5831
1.9003 | -93.6566
1.904 | -93.4832
1.9077 | -93.1158
1.9114 | -92.7883
1.9151 | -92.7226
1.9188 | -92.9887
1.9225 | -93.4175
1.9262 | -93.7401
1.9299 | -93.8741
1.9337 | -93.9591
1.9374 | -94.0191
1.9411 | -93.9129
1.9448 | -93.5803
1.9485 | -93.1462
1.9522 | -92.7951
1.9559 | -92.6628
1.9596 | -92.7544
1.9633 | -92.9548
1.9671 | -93.1621
1.9708 | -93.396
1.9745 | -93.6923
1.9782 | -93.965
1.9819 | -94.04
1.9856 | -93.8567
1.9893 | -93.5597
1.993 | -93.3188
1.9968 | -93.1667
2.0005 | -93.0823
2.0042 | -93.0282
2.0079 | -92.9927
2.0116 | -92.9938
2.0153 | -93.0522
2.019 | -93.1607
2.0227 | -93.2783
2.0264 | -93.3453
2.0302 | -93.3279
2.0339 | -93.2311
2.0376 | -93.1587
2.0413 | -93.2214
2.045 | -93.445
2.0487 | -93.7203
2.0524 | -93.8696
2.0561 | -93.8319
2.0598 | -93.7206
2.0636 | -93.671
2.0673 | -93.7624
2.071 | -94.0031
2.0747 | -94.31
2.0784 | -94.5587
2.0821 | -94.6924
2.0858 | -94.7428
2.0895 | -94.766
2.0932 | -94.7477
2.097 | -94.6568
2.1007 | -94.531
2.1044 | -94.4275
2.1081 | -94.3953
2.1118 | -94.3857
2.1155 | -94.344
2.1192 | -94.2671
2.1229 | -94.1757
2.1267 | -94.0647
2.1304 | -93.9612
2.1341 | -93.8693
2.1378 | -93.817
2.1415 | -93.8528
2.1452 | -93.8741
2.1489 | -93.8995
2.1526 | -93.9276
2.1563 | -93.9689
2.1601 | -94.0225
2.1638 | -94.0943
2.1675 | -94.1579
2.1712 | -94.2348
2.1749 | -94.3253
2.1786 | -94.4027
2.1823 | -94.4342
2.186 | -94.4193
2.1897 | -94.3892
2.1935 | -94.3591
2.1972 | -94.3382
2.2009 | -94.3214
2.2046 | -94.3104
2.2083 | -94.3456
2.212 | -94.4328
2.2157 | -94.5359
2.2194 | -94.5924
2.2231 | -94.5451
2.2269 | -94.3944
2.2306 | -94.201
2.2343 | -94.037
2.238 | -93.9074
2.2417 | -93.7993
2.2454 | -93.7082
2.2491 | -93.6556
2.2528 | -93.6507
2.2566 | -93.6699
2.2603 | -93.695
2.264 | -93.7403
2.2677 | -93.8378
2.2714 | -93.9851
2.2751 | -94.1308
2.2788 | -94.2032
2.2825 | -94.1511
2.2862 | -94.0265
2.29 | -93.9268
2.2937 | -93.9409
2.2974 | -94.0816
2.3011 | -94.3006
2.3048 | -94.5222
2.3085 | -94.6822
2.3122 | -94.7669
2.3159 | -94.8023
2.3196 | -94.8178
2.3234 | -94.8081
2.3271 | -94.7616
2.3308 | -94.6671
2.3345 | -94.5432
2.3382 | -94.4069
2.3419 | -94.2849
2.3456 | -94.2375
2.3493 | -94.3255
2.3531 | -94.5642
2.3568 | -94.8906
2.3605 | -95.202
2.3642 | -95.3933
2.3679 | -95.4052
2.3716 | -95.2813
2.3753 | -95.0701
2.379 | -94.8304
2.3827 | -94.6345
2.3865 | -94.5225
2.3902 | -94.5201
2.3939 | -94.605
2.3976 | -94.7481
2.4013 | -94.9062
2.405 | -95.0569
2.4087 | -95.1757
2.4124 | -95.2452
2.4161 | -95.2667
2.4199 | -95.2594
2.4236 | -95.2348
2.4273 | -95.2194
2.431 | -95.1998
2.4347 | -95.1792
2.4384 | -95.1605
2.4421 | -95.1842
2.4458 | -95.2391
2.4495 | -95.284
2.4533 | -95.267
2.457 | -95.2062
2.4607 | -95.154
2.4644 | -95.1378
2.4681 | -95.1422
2.4718 | -95.1499
2.4755 | -95.1703
2.4792 | -95.2337
2.483 | -95.3657
2.4867 | -95.5593
2.4904 | -95.7377
2.4941 | -95.8335
2.4978 | -95.8181
2.5015 | -95.71
2.5052 | -95.5909
2.5089 | -95.522
2.5126 | -95.5525
2.5164 | -95.6768
2.5201 | -95.8558
2.5238 | -96.0107
2.5275 | -96.0945
2.5312 | -96.1159
2.5349 | -96.0912
2.5386 | -96.0575
2.5423 | -96.0387
2.546 | -96.0435
2.5498 | -96.0833
2.5535 | -96.1045
2.5572 | -96.0427
2.5609 | -95.8569
2.5646 | -95.5935
2.5683 | -95.3404
2.572 | -95.1609
2.5757 | -95.0756
2.5794 | -95.0744
2.5832 | -95.1364
2.5869 | -95.2323
2.5906 | -95.3644
2.5943 | -95.5434
2.598 | -95.7623
2.6017 | -95.9656
2.6054 | -96.0893
2.6091 | -96.1034
2.6129 | -96.0144
2.6166 | -95.8876
2.6203 | -95.8028
2.624 | -95.8118
2.6277 | -95.9235
2.6314 | -96.0672
2.6351 | -96.1645
2.6388 | -96.1809
2.6425 | -96.1423
2.6463 | -96.0736
2.65 | -96.0189
2.6537 | -96.0211
2.6574 | -96.0619
2.6611 | -96.0998
2.6648 | -96.0906
2.6685 | -96.0122
2.6722 | -95.9101
2.6759 | -95.8648
2.6797 | -95.8231
2.6834 | -95.7965
2.6871 | -95.7607
2.6908 | -95.7219
2.6945 | -95.6773
2.6982 | -95.6309
2.7019 | -95.5803
2.7056 | -95.52
2.7093 | -95.4637
2.7131 | -95.4305
2.7168 | -95.4353
2.7205 | -95.4707
2.7242 | -95.5198
2.7279 | -95.556
2.7316 | -95.5724
2.7353 | -95.5688
2.739 | -95.545
2.7428 | -95.5172
2.7465 | -95.4952
2.7502 | -95.5086
2.7539 | -95.5716
2.7576 | -95.6928
2.7613 | -95.8466
2.765 | -96.0011
2.7687 | -96.1168
2.7724 | -96.1651
2.7762 | -96.1433
2.7799 | -96.0752
2.7836 | -95.9915
2.7873 | -95.903
2.791 | -95.8309
2.7947 | -95.7964
2.7984 | -95.803
2.8021 | -95.8497
2.8058 | -95.9215
2.8096 | -96.0084
2.8133 | -96.0931
2.817 | -96.1635
2.8207 | -96.2249
2.8244 | -96.2512
2.8281 | -96.2271
2.8318 | -96.1453
2.8355 | -96.0107
2.8392 | -95.8328
2.843 | -95.6584
2.8467 | -95.5354
2.8504 | -95.4799
2.8541 | -95.4933
2.8578 | -95.5697
2.8615 | -95.6764
2.8652 | -95.7792
2.8689 | -95.8393
2.8727 | -95.8211
2.8764 | -95.7208
2.8801 | -95.5877
2.8838 | -95.4984
2.8875 | -95.5187
2.8912 | -95.6455
2.8949 | -95.8539
2.8986 | -96.0862
2.9023 | -96.275
2.9061 | -96.3698
2.9098 | -96.371
2.9135 | -96.322
2.9172 | -96.242
2.9209 | -96.1764
2.9246 | -96.1532
2.9283 | -96.1855
2.932 | -96.2624
2.9357 | -96.3839
2.9395 | -96.5413
2.9432 | -96.7053
2.9469 | -96.86
2.9506 | -96.9824
2.9543 | -97.0629
2.958 | -97.0958
2.9617 | -97.0661
2.9654 | -96.9891
2.9691 | -96.8791
2.9729 | -96.7434
2.9766 | -96.6171
2.9803 | -96.5361
2.984 | -96.5111
2.9877 | -96.5511
2.9914 | -96.6526
2.9951 | -96.7909
2.9988 | -96.9213
3.0026 | -97.003
3.0063 | -97.0028
3.01 | -96.912
3.0137 | -96.7523
3.0174 | -96.5436
3.0211 | -96.352
3.0248 | -96.2328
3.0285 | -96.2239
3.0322 | -96.3392
3.036 | -96.5586
3.0397 | -96.8096
3.0434 | -97.0354
3.0471 | -97.184
3.0508 | -97.2037
3.0545 | -97.0924
3.0582 | -96.892
3.0619 | -96.6642
3.0656 | -96.4797
3.0694 | -96.3739
3.0731 | -96.369
3.0768 | -96.4364
3.0805 | -96.5102
3.0842 | -96.5452
3.0879 | -96.5303
3.0916 | -96.4542
3.0953 | -96.3547
3.099 | -96.2816
3.1028 | -96.2665
3.1065 | -96.3172
3.1102 | -96.4388
3.1139 | -96.6102
3.1176 | -96.8023
3.1213 | -96.9954
3.125 | -97.1792
3.1287 | -97.3279
3.1325 | -97.402
3.1362 | -97.379
3.1399 | -97.2937
3.1436 | -97.1653
3.1473 | -97.0238
3.151 | -96.9007
3.1547 | -96.836
3.1584 | -96.8459
3.1621 | -96.926
3.1659 | -97.0446
3.1696 | -97.1668
3.1733 | -97.2746
3.177 | -97.3575
3.1807 | -97.3877
3.1844 | -97.3851
3.1881 | -97.346
3.1918 | -97.2772
3.1955 | -97.1696
3.1993 | -97.0332
3.203 | -96.8964
3.2067 | -96.7747
3.2104 | -96.6917
3.2141 | -96.6153
3.2178 | -96.5759
3.2215 | -96.5719
3.2252 | -96.6142
3.2289 | -96.718
3.2327 | -96.8807
3.2364 | -97.0704
3.2401 | -97.2551
3.2438 | -97.4085
3.2475 | -97.5134
3.2512 | -97.5461
3.2549 | -97.5193
3.2586 | -97.4777
3.2624 | -97.4398
3.2661 | -97.4186
3.2698 | -97.42
3.2735 | -97.441
3.2772 | -97.4807
3.2809 | -97.5232
3.2846 | -97.5702
3.2883 | -97.6179
3.292 | -97.6541
3.2958 | -97.6788
3.2995 | -97.6813
3.3032 | -97.6684
3.3069 | -97.6386
3.3106 | -97.6019
3.3143 | -97.5589
3.318 | -97.5265
3.3217 | -97.4984
3.3254 | -97.4658
3.3292 | -97.4279
3.3329 | -97.3974
3.3366 | -97.3753
3.3403 | -97.359
3.344 | -97.3531
3.3477 | -97.3623
3.3514 | -97.3606
3.3551 | -97.3352
3.3588 | -97.2784
3.3626 | -97.2024
3.3663 | -97.1098
3.37 | -97.0197
3.3737 | -96.9462
3.3774 | -96.9033
3.3811 | -96.8955
3.3848 | -96.9145
3.3885 | -96.9578
3.3923 | -97.0249
3.396 | -97.1049
3.3997 | -97.2013
3.4034 | -97.3071
3.4071 | -97.4243
3.4108 | -97.541
3.4145 | -97.629
3.4182 | -97.6775
3.4219 | -97.6967
3.4257 | -97.7018
3.4294 | -97.7034
3.4331 | -97.7119
3.4368 | -97.715
3.4405 | -97.6962
3.4442 | -97.6413
3.4479 | -97.5547
3.4516 | -97.46
3.4553 | -97.3823
3.4591 | -97.3398
3.4628 | -97.3436
3.4665 | -97.3873
3.4702 | -97.4629
3.4739 | -97.5496
3.4776 | -97.6241
3.4813 | -97.6764
3.485 | -97.7045
3.4888 | -97.7231
3.4925 | -97.7502
3.4962 | -97.801
3.4999 | -97.8759
3.5036 | -97.9777
3.5073 | -98.0842
3.511 | -98.1781
3.5147 | -98.24
3.5184 | -98.2416
3.5222 | -98.1754
3.5259 | -98.0557
3.5296 | -97.9101
3.5333 | -97.7585
3.537 | -97.6226
3.5407 | -97.5222
3.5444 | -97.4606
3.5481 | -97.4454
3.5518 | -97.479
3.5556 | -97.5589
3.5593 | -97.6614
3.563 | -97.7777
3.5667 | -97.879
3.5704 | -97.9448
3.5741 | -97.9763
3.5778 | -97.9839
3.5815 | -97.98
3.5852 | -97.9885
3.589 | -98.0349
3.5927 | -98.0973
3.5964 | -98.144
3.6001 | -98.16
3.6038 | -98.1515
3.6075 | -98.1262
3.6112 | -98.1175
3.6149 | -98.1518
3.6187 | -98.2288
3.6224 | -98.3461
3.6261 | -98.4927
3.6298 | -98.6544
3.6335 | -98.8061
3.6372 | -98.9421
3.6409 | -99.0315
3.6446 | -99.0556
3.6483 | -99.0311
3.6521 | -98.9689
3.6558 | -98.864
3.6595 | -98.7354
3.6632 | -98.6119
3.6669 | -98.5044
3.6706 | -98.4357
3.6743 | -98.41
3.678 | -98.4148
3.6817 | -98.4296
3.6855 | -98.415
3.6892 | -98.3453
3.6929 | -98.2239
3.6966 | -98.0886
3.7003 | -97.9611
3.704 | -97.8533
3.7077 | -97.776
3.7114 | -97.7352
3.7151 | -97.7254
3.7189 | -97.7347
3.7226 | -97.763
3.7263 | -97.8113
3.73 | -97.8711
3.7337 | -97.9293
3.7374 | -97.9707
3.7411 | -98.0052
3.7448 | -98.0395
3.7485 | -98.0856
3.7523 | -98.1388
3.756 | -98.1986
3.7597 | -98.2737
3.7634 | -98.3503
3.7671 | -98.4199
3.7708 | -98.4789
3.7745 | -98.5097
3.7782 | -98.5287
3.782 | -98.535
3.7857 | -98.5325
3.7894 | -98.5443
3.7931 | -98.5877
3.7968 | -98.6558
3.8005 | -98.7365
3.8042 | -98.8236
3.8079 | -98.8966
3.8116 | -98.9519
3.8154 | -98.9738
3.8191 | -98.9626
3.8228 | -98.9298
3.8265 | -98.8857
3.8302 | -98.8339
3.8339 | -98.7813
3.8376 | -98.7342
3.8413 | -98.6983
3.845 | -98.6881
3.8488 | -98.6861
3.8525 | -98.6773
3.8562 | -98.6478
3.8599 | -98.5927
3.8636 | -98.5236
3.8673 | -98.453
3.871 | -98.3995
3.8747 | -98.3675
3.8784 | -98.3619
3.8822 | -98.3731
3.8859 | -98.3997
3.8896 | -98.4311
3.8933 | -98.4663
3.897 | -98.5086
3.9007 | -98.5719
3.9044 | -98.6341
3.9081 | -98.6895
3.9119 | -98.7534
3.9156 | -98.823
3.9193 | -98.8883
3.923 | -98.95
3.9267 | -98.9862
3.9304 | -99.0034
3.9341 | -98.9993
3.9378 | -98.9731
3.9415 | -98.9364
3.9453 | -98.8864
3.949 | -98.8175
3.9527 | -98.7572
3.9564 | -98.6993
3.9601 | -98.6506
3.9638 | -98.6161
3.9675 | -98.6261
3.9712 | -98.6758
3.9749 | -98.7639
3.9787 | -98.8762
3.9824 | -98.9983
3.9861 | -99.1174
3.9898 | -99.2102
3.9935 | -99.2524
3.9972 | -99.2446
4.0009 | -99.2027
4.0046 | -99.1354
4.0084 | -99.0734
4.0121 | -99.027
4.0158 | -99.0019
4.0195 | -99.0009
4.0232 | -99.0142
4.0269 | -99.032
4.0306 | -99.0376
4.0343 | -99.0465
4.038 | -99.057
4.0418 | -99.0723
4.0455 | -99.091
4.0492 | -99.1022
4.0529 | -99.1118
4.0566 | -99.126
4.0603 | -99.1548
4.064 | -99.1787
4.0677 | -99.1966
4.0714 | -99.1955
4.0752 | -99.1765
4.0789 | -99.14
4.0826 | -99.088
4.0863 | -99.0345
4.09 | -98.9861
4.0937 | -98.9576
4.0974 | -98.9459
4.1011 | -98.9594
4.1048 | -98.9969
4.1086 | -99.0562
4.1123 | -99.1213
4.116 | -99.1829
4.1197 | -99.2213
4.1234 | -99.2335
4.1271 | -99.2336
4.1308 | -99.2191
4.1345 | -99.1947
4.1383 | -99.1601
4.142 | -99.1226
4.1457 | -99.0865
4.1494 | -99.044
4.1531 | -99.0034
4.1568 | -98.9708
4.1605 | -98.9448
4.1642 | -98.9271
4.1679 | -98.9236
4.1717 | -98.9393
4.1754 | -98.9857
4.1791 | -99.0585
4.1828 | -99.1524
4.1865 | -99.2548
4.1902 | -99.3603
4.1939 | -99.4452
4.1976 | -99.4986
4.2013 | -99.5154
4.2051 | -99.5035
4.2088 | -99.4773
4.2125 | -99.4419
4.2162 | -99.4036
4.2199 | -99.3684
4.2236 | -99.3476
4.2273 | -99.3306
4.231 | -99.3219
4.2347 | -99.3236
4.2385 | -99.3259
4.2422 | -99.3354
4.2459 | -99.3474
4.2496 | -99.3548
4.2533 | -99.3482
4.257 | -99.3356
4.2607 | -99.3159
4.2644 | -99.2916
4.2682 | -99.2651
4.2719 | -99.2494
4.2756 | -99.2505
4.2793 | -99.2643
4.283 | -99.3021
4.2867 | -99.3291
4.2904 | -99.3621
4.2941 | -99.3899
4.2978 | -99.4103
4.3016 | -99.4338
4.3053 | -99.4727
4.309 | -99.5147
4.3127 | -99.5733
4.3164 | -99.6414
4.3201 | -99.7096
4.3238 | -99.775
4.3275 | -99.8259
4.3312 | -99.8672
4.335 | -99.8938
4.3387 | -99.9049
4.3424 | -99.8936
4.3461 | -99.8669
4.3498 | -99.8276
4.3535 | -99.7884
4.3572 | -99.7487
4.3609 | -99.7111
4.3646 | -99.6951
4.3684 | -99.6899
4.3721 | -99.6854
4.3758 | -99.687
4.3795 | -99.6922
4.3832 | -99.6751
4.3869 | -99.6505
4.3906 | -99.6142
4.3943 | -99.5575
4.3981 | -99.486
4.4018 | -99.4116
4.4055 | -99.3491
4.4092 | -99.3142
4.4129 | -99.3171
4.4166 | -99.3576
4.4203 | -99.4357
4.424 | -99.5456
4.4277 | -99.6859
4.4315 | -99.8373
4.4352 | -99.9915
4.4389 | -100.1276
4.4426 | -100.2467
4.4463 | -100.3319
4.45 | -100.3879
4.4537 | -100.428
4.4574 | -100.441
4.4611 | -100.4466
4.4649 | -100.4567
4.4686 | -100.4747
4.4723 | -100.4923
4.476 | -100.4982
4.4797 | -100.4953
4.4834 | -100.4796
4.4871 | -100.4301
4.4908 | -100.3377
4.4945 | -100.2166
4.4983 | -100.0722
4.502 | -99.9362
4.5057 | -99.8105
4.5094 | -99.7173
4.5131 | -99.6561
4.5168 | -99.6331
4.5205 | -99.639
4.5242 | -99.6677
4.528 | -99.7102
4.5317 | -99.7576
4.5354 | -99.8047
4.5391 | -99.8535
4.5428 | -99.8992
4.5465 | -99.9457
4.5502 | -99.9671
4.5539 | -99.9721
4.5576 | -99.9722
4.5614 | -99.9519
4.5651 | -99.9155
4.5688 | -99.8627
4.5725 | -99.7972
4.5762 | -99.744
4.5799 | -99.7151
4.5836 | -99.7133
4.5873 | -99.7459
4.591 | -99.8104
4.5948 | -99.9012
4.5985 | -100.0147
4.6022 | -100.1343
4.6059 | -100.2438
4.6096 | -100.3183
4.6133 | -100.366
4.617 | -100.3754
4.6207 | -100.3514
4.6244 | -100.2873
4.6282 | -100.1892
4.6319 | -100.0954
4.6356 | -100.0118
4.6393 | -99.9618
4.643 | -99.9459
4.6467 | -99.9557
4.6504 | -99.985
4.6541 | -100.0227
4.6579 | -100.0483
4.6616 | -100.0599
4.6653 | -100.0468
4.669 | -100.0151
4.6727 | -99.9708
4.6764 | -99.928
4.6801 | -99.8824
4.6838 | -99.8336
4.6875 | -99.7981
4.6913 | -99.7833
4.695 | -99.7892
4.6987 | -99.8135
4.7024 | -99.8475
4.7061 | -99.8953
4.7098 | -99.9557
4.7135 | -100.0176
4.7172 | -100.0736
4.7209 | -100.1186
4.7247 | -100.1426
4.7284 | -100.1554
4.7321 | -100.1529
4.7358 | -100.1405
4.7395 | -100.1283
4.7432 | -100.1294
4.7469 | -100.1339
4.7506 | -100.1467
4.7543 | -100.1635
4.7581 | -100.1777
4.7618 | -100.1876
4.7655 | -100.2015
4.7692 | -100.2199
4.7729 | -100.2445
4.7766 | -100.2785
4.7803 | -100.3224
4.784 | -100.3803
4.7878 | -100.4488
4.7915 | -100.5201
4.7952 | -100.5828
4.7989 | -100.6177
4.8026 | -100.6313
4.8063 | -100.6123
4.81 | -100.5579
4.8137 | -100.4818
4.8174 | -100.3788
4.8212 | -100.289
4.8249 | -100.2025
4.8286 | -100.123
4.8323 | -100.0585
4.836 | -100.0149
4.8397 | -99.9897
4.8434 | -99.9955
4.8471 | -100.0292
4.8508 | -100.0829
4.8546 | -100.1572
4.8583 | -100.247
4.862 | -100.3447
4.8657 | -100.4441
4.8694 | -100.5417
4.8731 | -100.6297
4.8768 | -100.6997
4.8805 | -100.7585
4.8842 | -100.8055
4.888 | -100.8458
4.8917 | -100.8802
4.8954 | -100.9133
4.8991 | -100.9442
4.9028 | -100.9628
4.9065 | -100.9607
4.9102 | -100.9505
4.9139 | -100.9403
4.9177 | -100.9256
4.9214 | -100.9065
4.9251 | -100.892
4.9288 | -100.8833
4.9325 | -100.878
4.9362 | -100.8711
4.9399 | -100.8619
4.9436 | -100.8469
4.9473 | -100.8177
4.9511 | -100.7904
4.9548 | -100.7632
4.9585 | -100.7299
4.9622 | -100.6878
4.9659 | -100.646
4.9696 | -100.621
4.9733 | -100.6059
4.977 | -100.6128
4.9807 | -100.6346
4.9845 | -100.6722
4.9882 | -100.7236
4.9919 | -100.7798
4.9956 | -100.8285
4.9993 | -100.8588
5.003 | -100.8664
5.0067 | -100.8466
5.0104 | -100.7963
5.0141 | -100.7254
5.0179 | -100.6431
5.0216 | -100.5641
5.0253 | -100.498
5.029 | -100.4361
5.0327 | -100.3883
5.0364 | -100.3638
5.0401 | -100.3614
5.0438 | -100.3709
5.0476 | -100.3896
5.0513 | -100.4217
5.055 | -100.4568
5.0587 | -100.4945
5.0624 | -100.5315
5.0661 | -100.5666
5.0698 | -100.5964
5.0735 | -100.6283
5.0772 | -100.6614
5.081 | -100.6956
5.0847 | -100.7361
5.0884 | -100.7837
5.0921 | -100.8357
5.0958 | -100.8899
5.0995 | -100.9382
5.1032 | -100.9847
5.1069 | -101.0276
5.1106 | -101.0581
5.1144 | -101.0724
5.1181 | -101.0645
5.1218 | -101.0411
5.1255 | -101.0014
5.1292 | -100.9566
5.1329 | -100.9137
5.1366 | -100.8759
5.1403 | -100.8485
5.1441 | -100.8263
5.1478 | -100.8177
5.1515 | -100.8228
5.1552 | -100.8354
5.1589 | -100.8659
5.1626 | -100.9126
5.1663 | -100.9735
5.17 | -101.0443
5.1737 | -101.1197
5.1775 | -101.1972
5.1812 | -101.2585
5.1849 | -101.3018
5.1886 | -101.3323
5.1923 | -101.3345
5.196 | -101.3151
5.1997 | -101.2887
5.2034 | -101.2472
5.2071 | -101.1953
5.2109 | -101.1465
5.2146 | -101.1112
5.2183 | -101.0853
5.222 | -101.0743
5.2257 | -101.075
5.2294 | -101.094
5.2331 | -101.1232
5.2368 | -101.1525
5.2405 | -101.1898
5.2443 | -101.2328
5.248 | -101.2633
5.2517 | -101.279
5.2554 | -101.2713
5.2591 | -101.2457
5.2628 | -101.1997
5.2665 | -101.1404
5.2702 | -101.0866
5.274 | -101.0304
5.2777 | -100.9718
5.2814 | -100.9195
5.2851 | -100.873
5.2888 | -100.8408
5.2925 | -100.819
5.2962 | -100.8015
5.2999 | -100.7975
5.3036 | -100.8033
5.3074 | -100.8065
5.3111 | -100.8112
5.3148 | -100.8088
5.3185 | -100.7958
5.3222 | -100.7807
5.3259 | -100.7724
5.3296 | -100.7738
5.3333 | -100.7942
5.337 | -100.8321
5.3408 | -100.8786
5.3445 | -100.9358
5.3482 | -101.0021
5.3519 | -101.0778
5.3556 | -101.1371
5.3593 | -101.1929
5.363 | -101.2513
5.3667 | -101.3001
5.3704 | -101.3215
5.3742 | -101.3366
5.3779 | -101.3501
5.3816 | -101.3482
5.3853 | -101.3297
5.389 | -101.302
5.3927 | -101.2758
5.3964 | -101.2491
5.4001 | -101.2217
5.4039 | -101.1868
5.4076 | -101.1551
5.4113 | -101.1217
5.415 | -101.0781
5.4187 | -101.0376
5.4224 | -101.0068
5.4261 | -100.9817
5.4298 | -100.9648
5.4335 | -100.9592
5.4373 | -100.9679
5.441 | -100.9941
5.4447 | -101.0303
5.4484 | -101.0722
5.4521 | -101.1225
5.4558 | -101.1801
5.4595 | -101.2267
5.4632 | -101.2635
5.4669 | -101.2886
5.4707 | -101.3115
5.4744 | -101.3265
5.4781 | -101.3305
5.4818 | -101.3335
5.4855 | -101.3337
5.4892 | -101.3417
5.4929 | -101.3432
5.4966 | -101.3353
5.5003 | -101.3276
5.5041 | -101.31
5.5078 | -101.2744
5.5115 | -101.2241
5.5152 | -101.1612
5.5189 | -101.0868
5.5226 | -101.0112
5.5263 | -100.9353
5.53 | -100.8665
5.5338 | -100.8111
5.5375 | -100.7811
5.5412 | -100.7731
5.5449 | -100.7923
5.5486 | -100.8373
5.5523 | -100.9006
5.556 | -100.9733
5.5597 | -101.0488
5.5634 | -101.1225
5.5672 | -101.1994
5.5709 | -101.2702
5.5746 | -101.3375
5.5783 | -101.4074
5.582 | -101.4825
5.5857 | -101.5536
5.5894 | -101.6159
5.5931 | -101.6796
5.5968 | -101.7321
5.6006 | -101.7699
5.6043 | -101.7935
5.608 | -101.8045
5.6117 | -101.8073
5.6154 | -101.8042
5.6191 | -101.8007
5.6228 | -101.7855
5.6265 | -101.7729
5.6302 | -101.7645
5.634 | -101.7557
5.6377 | -101.7437
5.6414 | -101.72
5.6451 | -101.6937
5.6488 | -101.6601
5.6525 | -101.6101
5.6562 | -101.5586
5.6599 | -101.4945
5.6637 | -101.4257
5.6674 | -101.3553
5.6711 | -101.2931
5.6748 | -101.2427
5.6785 | -101.209
5.6822 | -101.1955
5.6859 | -101.2027
5.6896 | -101.2286
5.6933 | -101.2616
5.6971 | -101.3109
5.7008 | -101.3705
5.7045 | -101.4255
5.7082 | -101.4767
5.7119 | -101.517
5.7156 | -101.532
5.7193 | -101.5296
5.723 | -101.5163
5.7267 | -101.4878
5.7305 | -101.45
5.7342 | -101.4052
5.7379 | -101.3575
5.7416 | -101.3134
5.7453 | -101.2763
5.749 | -101.2388
5.7527 | -101.203
5.7564 | -101.1694
5.7601 | -101.145
5.7639 | -101.1308
5.7676 | -101.1341
5.7713 | -101.1567
5.775 | -101.1958
5.7787 | -101.2544
5.7824 | -101.317
5.7861 | -101.3848
5.7898 | -101.4584
5.7936 | -101.5263
5.7973 | -101.5862
5.801 | -101.6435
5.8047 | -101.696
5.8084 | -101.7341
5.8121 | -101.7581
5.8158 | -101.7638
5.8195 | -101.7593
5.8232 | -101.7499
5.827 | -101.7322
5.8307 | -101.7058
5.8344 | -101.6782
5.8381 | -101.6572
5.8418 | -101.6376
5.8455 | -101.6238
5.8492 | -101.6254
5.8529 | -101.6407
5.8566 | -101.6746
5.8604 | -101.7283
5.8641 | -101.7949
5.8678 | -101.8726
5.8715 | -101.9574
5.8752 | -102.0371
5.8789 | -102.1062
5.8826 | -102.1574
5.8863 | -102.1353
5.89 | -102.1374
5.8938 | -102.1229
5.8975 | -102.0997
5.9012 | -102.0707
5.9049 | -102.0351
5.9086 | -101.9947
5.9123 | -101.9492
5.916 | -101.9099
5.9197 | -101.8841
5.9235 | -101.8666
5.9272 | -101.8503
5.9309 | -101.8429
5.9346 | -101.8424
5.9383 | -101.8331
5.942 | -101.8342
5.9457 | -101.8375
5.9494 | -101.842
5.9531 | -101.8468
5.9569 | -101.8511
5.9606 | -101.8502
5.9643 | -101.8556
5.968 | -101.8632
5.9717 | -101.8752
5.9754 | -101.8934
5.9791 | -101.9065
5.9828 | -101.9275
5.9865 | -101.9542
5.9903 | -101.9863
5.994 | -102.0197
5.9977 | -102.0472
6.0014 | -102.0646
6.0051 | -102.0734
6.0088 | -102.0671
6.0125 | -102.0464
6.0162 | -102.0174
6.0199 | -101.9821
6.0237 | -101.9397
6.0274 | -101.8833
6.0311 | -101.8289
6.0348 | -101.7647
6.0385 | -101.698
6.0422 | -101.642
6.0459 | -101.5849
6.0496 | -101.5325
6.0534 | -101.4902
6.0571 | -101.451
6.0608 | -101.4166
6.0645 | -101.3928
6.0682 | -101.374
6.0719 | -101.3611
6.0756 | -101.3538
6.0793 | -101.3561
6.083 | -101.3619
6.0868 | -101.3722
6.0905 | -101.3953
6.0942 | -101.4277
6.0979 | -101.4656
6.1016 | -101.5036
6.1053 | -101.5473
6.109 | -101.5995
6.1127 | -101.6568
6.1164 | -101.727
6.1202 | -101.8101
6.1239 | -101.893
6.1276 | -101.9778
6.1313 | -102.0616
6.135 | -102.1325
6.1387 | -102.2038
6.1424 | -102.2581
6.1461 | -102.2896
6.1498 | -102.3037
6.1536 | -102.299
6.1573 | -102.278
6.161 | -102.2589
6.1647 | -102.2389
6.1684 | -102.2191
6.1721 | -102.2112
6.1758 | -102.2182
6.1795 | -102.2457
6.1833 | -102.2798
6.187 | -102.3188
6.1907 | -102.3631
6.1944 | -102.4109
6.1981 | -102.4516
6.2018 | -102.4771
6.2055 | -102.4942
6.2092 | -102.4986
6.2129 | -102.491
6.2167 | -102.4723
6.2204 | -102.4429
6.2241 | -102.412
6.2278 | -102.378
6.2315 | -102.3412
6.2352 | -102.3127
6.2389 | -102.2948
6.2426 | -102.2875
6.2463 | -102.2813
6.2501 | -102.2726
6.2538 | -102.2699
6.2575 | -102.28
6.2612 | -102.2945
6.2649 | -102.3161
6.2686 | -102.3513
6.2723 | -102.3894
6.276 | -102.4298
6.2797 | -102.4824
6.2835 | -102.537
6.2872 | -102.5934
6.2909 | -102.6548
6.2946 | -102.7157
6.2983 | -102.7715
6.302 | -102.8181
6.3057 | -102.8539
6.3094 | -102.875
6.3132 | -102.8758
6.3169 | -102.8547
6.3206 | -102.8141
6.3243 | -102.7519
6.328 | -102.664
6.3317 | -102.5641
6.3354 | -102.4561
6.3391 | -102.3492
6.3428 | -102.2492
6.3466 | -102.1488
6.3503 | -102.0544
6.354 | -101.9816
6.3577 | -101.9295
6.3614 | -101.8985
6.3651 | -101.8899
6.3688 | -101.9058
6.3725 | -101.9346
6.3762 | -101.9826
6.38 | -102.0403
6.3837 | -102.1091
6.3874 | -102.1792
6.3911 | -102.2516
6.3948 | -102.3199
6.3985 | -102.3803
6.4022 | -102.4289
6.4059 | -102.4644
6.4097 | -102.4924
6.4134 | -102.5126
6.4171 | -102.5309
6.4208 | -102.5494
6.4245 | -102.5534
6.4282 | -102.5637
6.4319 | -102.5747
6.4356 | -102.5895
6.4393 | -102.601
6.4431 | -102.6158
6.4468 | -102.633
6.4505 | -102.6495
6.4542 | -102.6551
6.4579 | -102.6576
6.4616 | -102.6567
6.4653 | -102.6509
6.469 | -102.643
6.4727 | -102.6239
6.4765 | -102.6011
6.4802 | -102.5743
6.4839 | -102.5523
6.4876 | -102.5232
6.4913 | -102.4952
6.495 | -102.4765
6.4987 | -102.4666
6.5024 | -102.4621
6.5061 | -102.4632
6.5099 | -102.4723
6.5136 | -102.4837
6.5173 | -102.5048
6.521 | -102.5329
6.5247 | -102.5655
6.5284 | -102.5926
6.5321 | -102.6102
6.5358 | -102.6277
6.5396 | -102.648
6.5433 | -102.6611
6.547 | -102.6641
6.5507 | -102.6585
6.5544 | -102.6584
6.5581 | -102.659
6.5618 | -102.6589
6.5655 | -102.6548
6.5692 | -102.655
6.573 | -102.6639
6.5767 | -102.6793
6.5804 | -102.6941
6.5841 | -102.7118
6.5878 | -102.7295
6.5915 | -102.7399
6.5952 | -102.7381
6.5989 | -102.7326
6.6026 | -102.7185
6.6064 | -102.6926
6.6101 | -102.6611
6.6138 | -102.6317
6.6175 | -102.5931
6.6212 | -102.56
6.6249 | -102.5264
6.6286 | -102.4861
6.6323 | -102.4557
6.636 | -102.4228
6.6398 | -102.3913
6.6435 | -102.3643
6.6472 | -102.342
6.6509 | -102.3253
6.6546 | -102.3143
6.6583 | -102.3088
6.662 | -102.3146
6.6657 | -102.3335
6.6695 | -102.3612
6.6732 | -102.4017
6.6769 | -102.4489
6.6806 | -102.5015
6.6843 | -102.5567
6.688 | -102.6138
6.6917 | -102.6673
6.6954 | -102.7089
6.6991 | -102.7352
6.7029 | -102.7599
6.7066 | -102.7774
6.7103 | -102.7869
6.714 | -102.7901
6.7177 | -102.792
6.7214 | -102.789
6.7251 | -102.7887
6.7288 | -102.8008
6.7325 | -102.8108
6.7363 | -102.8201
6.74 | -102.83
6.7437 | -102.8442
6.7474 | -102.8611
6.7511 | -102.8744
6.7548 | -102.8831
6.7585 | -102.8941
6.7622 | -102.8992
6.7659 | -102.8982
6.7697 | -102.8903
6.7734 | -102.8718
6.7771 | -102.8512
6.7808 | -102.8263
6.7845 | -102.7934
6.7882 | -102.7542
6.7919 | -102.7117
6.7956 | -102.6736
6.7993 | -102.6384
6.8031 | -102.6073
6.8068 | -102.5876
6.8105 | -102.5789
6.8142 | -102.581
6.8179 | -102.5916
6.8216 | -102.6078
6.8253 | -102.6266
6.829 | -102.6466
6.8328 | -102.6665
6.8365 | -102.6813
6.8402 | -102.694
6.8439 | -102.6998
6.8476 | -102.6929
6.8513 | -102.6768
6.855 | -102.6521
6.8587 | -102.6159
6.8624 | -102.5801
6.8662 | -102.5499
6.8699 | -102.521
6.8736 | -102.4951
6.8773 | -102.4768
6.881 | -102.4721
6.8847 | -102.4753
6.8884 | -102.4913
6.8921 | -102.5202
6.8958 | -102.5499
6.8996 | -102.5845
6.9033 | -102.6179
6.907 | -102.6519
6.9107 | -102.6883
6.9144 | -102.7213
6.9181 | -102.7481
6.9218 | -102.7699
6.9255 | -102.7833
6.9293 | -102.7929
6.933 | -102.795
6.9367 | -102.7939
6.9404 | -102.779
6.9441 | -102.7591
6.9478 | -102.7346
6.9515 | -102.7012
6.9552 | -102.6612
6.9589 | -102.6068
6.9627 | -102.5679
6.9664 | -102.5259
6.9701 | -102.4812
6.9738 | -102.4414
6.9775 | -102.4125
6.9812 | -102.3877
6.9849 | -102.3695
6.9886 | -102.3645
6.9923 | -102.3635
6.9961 | -102.3698
6.9998 | -102.3885
7.0035 | -102.4186
7.0072 | -102.4588
7.0109 | -102.5055
7.0146 | -102.5551
7.0183 | -102.6108
7.022 | -102.6712
7.0257 | -102.7292
7.0295 | -102.7858
7.0332 | -102.8386
7.0369 | -102.8877
7.0406 | -102.9273
7.0443 | -102.9617
7.048 | -102.9826
7.0517 | -102.9917
7.0554 | -102.993
7.0592 | -102.9804
7.0629 | -102.961
7.0666 | -102.9344
7.0703 | -102.8989
7.074 | -102.861
7.0777 | -102.8219
7.0814 | -102.7857
7.0851 | -102.7489
7.0888 | -102.7179
7.0926 | -102.6925
7.0963 | -102.6753
7.1 | -102.6648
7.1037 | -102.6593
7.1074 | -102.6592
7.1111 | -102.6561
7.1148 | -102.6573
7.1185 | -102.6622
7.1222 | -102.673
7.126 | -102.6864
7.1297 | -102.7
7.1334 | -102.7165
7.1371 | -102.7294
7.1408 | -102.738
7.1445 | -102.7582
7.1482 | -102.7851
7.1519 | -102.8146
7.1556 | -102.852
7.1594 | -102.8874
7.1631 | -102.9195
7.1668 | -102.9581
7.1705 | -102.9926
7.1742 | -103.0226
7.1779 | -103.0493
7.1816 | -103.0743
7.1853 | -103.0974
7.1891 | -103.1094
7.1928 | -103.117
7.1965 | -103.1237
7.2002 | -103.1255
7.2039 | -103.121
7.2076 | -103.1111
7.2113 | -103.0997
7.215 | -103.0899
7.2187 | -103.0833
7.2225 | -103.0731
7.2262 | -103.0641
7.2299 | -103.0546
7.2336 | -103.043
7.2373 | -103.0339
7.241 | -103.0241
7.2447 | -103.0186
7.2484 | -103.0204
7.2521 | -103.0178
7.2559 | -103.0211
7.2596 | -103.028
7.2633 | -103.0343
7.267 | -103.0526
7.2707 | -103.0721
7.2744 | -103.0855
7.2781 | -103.1064
7.2818 | -103.1368
7.2855 | -103.1681
7.2893 | -103.1974
7.293 | -103.2222
7.2967 | -103.2509
7.3004 | -103.2764
7.3041 | -103.2957
7.3078 | -103.3075
7.3115 | -103.3152
7.3152 | -103.3199
7.319 | -103.3115
7.3227 | -103.2905
7.3264 | -103.2674
7.3301 | -103.2451
7.3338 | -103.2189
7.3375 | -103.1951
7.3412 | -103.1718
7.3449 | -103.1538
7.3486 | -103.1376
7.3524 | -103.1217
7.3561 | -103.1026
7.3598 | -103.0878
7.3635 | -103.08
7.3672 | -103.0726
7.3709 | -103.0681
7.3746 | -103.0675
7.3783 | -103.0708
7.382 | -103.0803
7.3858 | -103.0854
7.3895 | -103.0926
7.3932 | -103.1081
7.3969 | -103.1301
7.4006 | -103.1538
7.4043 | -103.1793
7.408 | -103.2076
7.4117 | -103.2408
7.4154 | -103.2785
7.4192 | -103.3213
7.4229 | -103.3604
7.4266 | -103.4029
7.4303 | -103.4496
7.434 | -103.4891
7.4377 | -103.5174
7.4414 | -103.5383
7.4451 | -103.5511
7.4489 | -103.5539
7.4526 | -103.5371
7.4563 | -103.5116
7.46 | -103.4808
7.4637 | -103.4469
7.4674 | -103.4124
7.4711 | -103.3732
7.4748 | -103.3403
7.4785 | -103.3088
7.4823 | -103.2701
7.486 | -103.2385
7.4897 | -103.2071
7.4934 | -103.1933
7.4971 | -103.1751
7.5008 | -103.1577
7.5045 | -103.1428
7.5082 | -103.1275
7.5119 | -103.1206
7.5157 | -103.1184
7.5194 | -103.1204
7.5231 | -103.1177
7.5268 | -103.1143
7.5305 | -103.1051
7.5342 | -103.0934
7.5379 | -103.0786
7.5416 | -103.0594
7.5453 | -103.0395
7.5491 | -103.0193
7.5528 | -103.003
7.5565 | -102.987
7.5602 | -102.9804
7.5639 | -102.9842
7.5676 | -102.9939
7.5713 | -103.011
7.575 | -103.0349
7.5788 | -103.0634
7.5825 | -103.0972
7.5862 | -103.1331
7.5899 | -103.1682
7.5936 | -103.2075
7.5973 | -103.2494
7.601 | -103.2881
7.6047 | -103.3171
7.6084 | -103.3385
7.6122 | -103.3587
7.6159 | -103.3741
7.6196 | -103.3858
7.6233 | -103.3976
7.627 | -103.4051
7.6307 | -103.4092
7.6344 | -103.4159
7.6381 | -103.4224
7.6418 | -103.4354
7.6456 | -103.4517
7.6493 | -103.4709
7.653 | -103.4905
7.6567 | -103.5125
7.6604 | -103.5407
7.6641 | -103.5755
7.6678 | -103.6067
7.6715 | -103.64
7.6752 | -103.671
7.679 | -103.6954
7.6827 | -103.7182
7.6864 | -103.7362
7.6901 | -103.7427
7.6938 | -103.7448
7.6975 | -103.7431
7.7012 | -103.7364
7.7049 | -103.7254
7.7087 | -103.7092
7.7124 | -103.6877
7.7161 | -103.6659
7.7198 | -103.6482
7.7235 | -103.6293
7.7272 | -103.6028
7.7309 | -103.5814
7.7346 | -103.5622
7.7383 | -103.5375
7.7421 | -103.5182
7.7458 | -103.505
7.7495 | -103.4922
7.7532 | -103.4883
7.7569 | -103.4834
7.7606 | -103.4838
7.7643 | -103.4921
7.768 | -103.5024
7.7717 | -103.5179
7.7755 | -103.5371
7.7792 | -103.5589
7.7829 | -103.5844
7.7866 | -103.6071
7.7903 | -103.6318
7.794 | -103.6571
7.7977 | -103.6786
7.8014 | -103.6947
7.8052 | -103.7137
7.8089 | -103.7236
7.8126 | -103.723
7.8163 | -103.7122
7.82 | -103.6921
7.8237 | -103.66
7.8274 | -103.6187
7.8311 | -103.579
7.8348 | -103.5343
7.8386 | -103.4868
7.8423 | -103.4449
7.846 | -103.3976
7.8497 | -103.3544
7.8534 | -103.3249
7.8571 | -103.3013
7.8608 | -103.2872
7.8645 | -103.2845
7.8682 | -103.2927
7.872 | -103.3125
7.8757 | -103.3395
7.8794 | -103.3681
7.8831 | -103.4079
7.8868 | -103.4537
7.8905 | -103.5034
7.8942 | -103.5563
7.8979 | -103.6125
7.9016 | -103.6702
7.9054 | -103.7246
7.9091 | -103.7772
7.9128 | -103.8255
7.9165 | -103.8707
7.9202 | -103.9016
7.9239 | -103.9259
7.9276 | -103.942
7.9313 | -103.9506
7.9351 | -103.9513
7.9388 | -103.9432
7.9425 | -103.9189
7.9462 | -103.8842
7.9499 | -103.8521
7.9536 | -103.8139
7.9573 | -103.7688
7.961 | -103.728
7.9647 | -103.6849
7.9685 | -103.636
7.9722 | -103.5846
7.9759 | -103.5362
7.9796 | -103.4979
7.9833 | -103.4644
7.987 | -103.428
7.9907 | -103.3959
7.9944 | -103.3738
7.9981 | -103.3542
8.0019 | -103.3334
8.0056 | -103.3172
8.0093 | -103.3062
8.013 | -103.3033
8.0167 | -103.303
8.0204 | -103.3043
8.0241 | -103.3088
8.0278 | -103.3126
8.0315 | -103.3283
8.0353 | -103.3417
8.039 | -103.3608
8.0427 | -103.3805
8.0464 | -103.4072
8.0501 | -103.4361
8.0538 | -103.4687
8.0575 | -103.4995
8.0612 | -103.5377
8.0649 | -103.576
8.0687 | -103.6049
8.0724 | -103.6367
8.0761 | -103.6661
8.0798 | -103.6923
8.0835 | -103.7173
8.0872 | -103.7393
8.0909 | -103.7598
8.0946 | -103.7718
8.0984 | -103.7768
8.1021 | -103.7801
8.1058 | -103.786
8.1095 | -103.7847
8.1132 | -103.7771
8.1169 | -103.7709
8.1206 | -103.7694
8.1243 | -103.7609
8.128 | -103.7514
8.1318 | -103.7406
8.1355 | -103.7349
8.1392 | -103.7319
8.1429 | -103.7251
8.1466 | -103.7263
8.1503 | -103.7332
8.154 | -103.7311
8.1577 | -103.7278
8.1614 | -103.7328
8.1652 | -103.7421
8.1689 | -103.7478
8.1726 | -103.7577
8.1763 | -103.7653
8.18 | -103.7801
8.1837 | -103.7935
8.1874 | -103.8067
8.1911 | -103.8183
8.1949 | -103.8285
8.1986 | -103.841
8.2023 | -103.8464
8.206 | -103.845
8.2097 | -103.8454
8.2134 | -103.8455
8.2171 | -103.8461
8.2208 | -103.846
8.2245 | -103.8488
8.2283 | -103.8508
8.232 | -103.8508
8.2357 | -103.8452
8.2394 | -103.8366
8.2431 | -103.8294
8.2468 | -103.8253
8.2505 | -103.8136
8.2542 | -103.8031
8.2579 | -103.7915
8.2617 | -103.7745
8.2654 | -103.7617
8.2691 | -103.7482
8.2728 | -103.7282
8.2765 | -103.7211
8.2802 | -103.7198
8.2839 | -103.7146
8.2876 | -103.7139
8.2913 | -103.7198
8.2951 | -103.7243
8.2988 | -103.7329
8.3025 | -103.74
8.3062 | -103.7495
8.3099 | -103.7595
8.3136 | -103.774
8.3173 | -103.7867
8.321 | -103.8005
8.3248 | -103.815
8.3285 | -103.8346
8.3322 | -103.8537
8.3359 | -103.8712
8.3396 | -103.8936
8.3433 | -103.9158
8.347 | -103.9354
8.3507 | -103.9563
8.3544 | -103.9808
8.3582 | -104.0011
8.3619 | -104.0127
8.3656 | -104.0207
8.3693 | -104.0251
8.373 | -104.0312
8.3767 | -104.0252
8.3804 | -104.0198
8.3841 | -104.0189
8.3878 | -104.016
8.3916 | -104.0091
8.3953 | -104.0114
8.399 | -104.0056
8.4027 | -103.9967
8.4064 | -103.9894
8.4101 | -103.9875
8.4138 | -103.986
8.4175 | -103.9817
8.4212 | -103.9735
8.425 | -103.9659
8.4287 | -103.9506
8.4324 | -103.9318
8.4361 | -103.9196
8.4398 | -103.9089
8.4435 | -103.8938
8.4472 | -103.8796
8.4509 | -103.8659
8.4547 | -103.8526
8.4584 | -103.834
8.4621 | -103.8087
8.4658 | -103.7941
8.4695 | -103.7805
8.4732 | -103.7643
8.4769 | -103.7469
8.4806 | -103.7325
8.4843 | -103.7156
8.4881 | -103.6978
8.4918 | -103.6838
8.4955 | -103.6678
8.4992 | -103.6583
8.5029 | -103.651
8.5066 | -103.6471
8.5103 | -103.6467
8.514 | -103.651
8.5177 | -103.6587
8.5215 | -103.6795
8.5252 | -103.7088
8.5289 | -103.7428
8.5326 | -103.7768
8.5363 | -103.8184
8.54 | -103.8678
8.5437 | -103.9211
8.5474 | -103.9725
8.5511 | -104.026
8.5549 | -104.0741
8.5586 | -104.1082
8.5623 | -104.1338
8.566 | -104.1281
8.5697 | -104.1458
8.5734 | -104.1535
8.5771 | -104.1441
8.5808 | -104.1314
8.5846 | -104.1188
8.5883 | -104.0984
8.592 | -104.0618
8.5957 | -104.0218
8.5994 | -103.9808
8.6031 | -103.9411
8.6068 | -103.9046
8.6105 | -103.871
8.6142 | -103.8411
8.618 | -103.8165
8.6217 | -103.8013
8.6254 | -103.7933
8.6291 | -103.7937
8.6328 | -103.801
8.6365 | -103.8201
8.6402 | -103.8487
8.6439 | -103.8831
8.6476 | -103.9211
8.6514 | -103.9662
8.6551 | -104.0153
8.6588 | -104.0662
8.6625 | -104.116
8.6662 | -104.1654
8.6699 | -104.2119
8.6736 | -104.2522
8.6773 | -104.2889
8.681 | -104.326
8.6848 | -104.3611
8.6885 | -104.3951
8.6922 | -104.4252
8.6959 | -104.4429
8.6996 | -104.4541
8.7033 | -104.4676
8.707 | -104.4833
8.7107 | -104.4885
8.7145 | -104.4926
8.7182 | -104.4972
8.7219 | -104.4955
8.7256 | -104.493
8.7293 | -104.4914
8.733 | -104.4894
8.7367 | -104.4863
8.7404 | -104.487
8.7441 | -104.4887
8.7479 | -104.4909
8.7516 | -104.4999
8.7553 | -104.5104
8.759 | -104.5197
8.7627 | -104.5314
8.7664 | -104.5401
8.7701 | -104.5489
8.7738 | -104.5617
8.7775 | -104.5704
8.7813 | -104.5687
8.785 | -104.572
8.7887 | -104.5679
8.7924 | -104.5636
8.7961 | -104.5561
8.7998 | -104.5406
8.8035 | -104.5227
8.8072 | -104.5049
8.8109 | -104.4813
8.8147 | -104.4598
8.8184 | -104.4376
8.8221 | -104.4101
8.8258 | -104.3806
8.8295 | -104.3562
8.8332 | -104.3365
8.8369 | -104.3197
8.8406 | -104.3052
8.8444 | -104.2934
8.8481 | -104.2888
8.8518 | -104.2899
8.8555 | -104.2914
8.8592 | -104.29
8.8629 | -104.2866
8.8666 | -104.282
8.8703 | -104.2778
8.874 | -104.2711
8.8778 | -104.2624
8.8815 | -104.2475
8.8852 | -104.2252
8.8889 | -104.202
8.8926 | -104.1756
8.8963 | -104.1445
8.9 | -104.1081
8.9037 | -104.0674
8.9074 | -104.0275
8.9112 | -103.9864
8.9149 | -103.9419
8.9186 | -103.9042
8.9223 | -103.8648
8.926 | -103.828
8.9297 | -103.8019
8.9334 | -103.7805
8.9371 | -103.7647
8.9408 | -103.7579
8.9446 | -103.7565
8.9483 | -103.7651
8.952 | -103.7884
8.9557 | -103.8235
8.9594 | -103.8646
8.9631 | -103.9052
8.9668 | -103.9518
8.9705 | -104.0037
8.9743 | -104.0602
8.978 | -104.1174
8.9817 | -104.1822
8.9854 | -104.2453
8.9891 | -104.3106
8.9928 | -104.368
8.9965 | -104.4198
9.0002 | -104.4621
9.0039 | -104.4974
9.0077 | -104.5237
9.0114 | -104.5419
9.0151 | -104.5548
9.0188 | -104.5606
9.0225 | -104.5609
9.0262 | -104.5576
9.0299 | -104.5499
9.0336 | -104.5365
9.0373 | -104.5308
9.0411 | -104.5283
9.0448 | -104.519
9.0485 | -104.5119
9.0522 | -104.5109
9.0559 | -104.5139
9.0596 | -104.5143
9.0633 | -104.5167
9.067 | -104.5192
9.0707 | -104.5194
9.0745 | -104.522
9.0782 | -104.5296
9.0819 | -104.5355
9.0856 | -104.5357
9.0893 | -104.5265
9.093 | -104.5204
9.0967 | -104.5089
9.1004 | -104.4836
9.1042 | -104.4695
9.1079 | -104.4541
9.1116 | -104.4365
9.1153 | -104.4086
9.119 | -104.385
9.1227 | -104.3601
9.1264 | -104.3369
9.1301 | -104.3177
9.1338 | -104.2977
9.1376 | -104.2749
9.1413 | -104.2528
9.145 | -104.2313
9.1487 | -104.2111
9.1524 | -104.1888
9.1561 | -104.1708
9.1598 | -104.1512
9.1635 | -104.134
9.1672 | -104.1171
9.171 | -104.0927
9.1747 | -104.0694
9.1784 | -104.0507
9.1821 | -104.0311
9.1858 | -104.0112
9.1895 | -103.9932
9.1932 | -103.9822
9.1969 | -103.9752
9.2006 | -103.97
9.2044 | -103.9723
9.2081 | -103.9798
9.2118 | -103.9874
9.2155 | -104.0025
9.2192 | -104.0302
9.2229 | -104.0638
9.2266 | -104.1002
9.2303 | -104.1374
9.2341 | -104.1713
9.2378 | -104.2037
9.2415 | -104.2372
9.2452 | -104.2645
9.2489 | -104.2874
9.2526 | -104.3122
9.2563 | -104.3252
9.26 | -104.3306
9.2637 | -104.3327
9.2675 | -104.3262
9.2712 | -104.3177
9.2749 | -104.3092
9.2786 | -104.2993
9.2823 | -104.2899
9.286 | -104.2802
9.2897 | -104.2724
9.2934 | -104.2644
9.2971 | -104.2629
9.3009 | -104.2645
9.3046 | -104.27
9.3083 | -104.2761
9.312 | -104.2835
9.3157 | -104.2926
9.3194 | -104.3013
9.3231 | -104.3119
9.3268 | -104.325
9.3306 | -104.3344
9.3343 | -104.3412
9.338 | -104.3471
9.3417 | -104.3495
9.3454 | -104.3499
9.3491 | -104.3509
9.3528 | -104.3479
9.3565 | -104.338
9.3602 | -104.3268
9.364 | -104.318
9.3677 | -104.3094
9.3714 | -104.2997
9.3751 | -104.2895
9.3788 | -104.2773
9.3825 | -104.2675
9.3862 | -104.2619
9.3899 | -104.2615
9.3936 | -104.2678
9.3974 | -104.27
9.4011 | -104.272
9.4048 | -104.2776
9.4085 | -104.2893
9.4122 | -104.3023
9.4159 | -104.3134
9.4196 | -104.3224
9.4233 | -104.3283
9.427 | -104.3331
9.4308 | -104.3379
9.4345 | -104.3368
9.4382 | -104.3305
9.4419 | -104.3227
9.4456 | -104.3189
9.4493 | -104.309
9.453 | -104.2946
9.4567 | -104.2785
9.4604 | -104.2618
9.4642 | -104.243
9.4679 | -104.2254
9.4716 | -104.2061
9.4753 | -104.1818
9.479 | -104.1558
9.4827 | -104.1272
9.4864 | -104.0985
9.4901 | -104.0711
9.4939 | -104.0392
9.4976 | -104.0084
9.5013 | -103.9693
9.505 | -103.9293
9.5087 | -103.8972
9.5124 | -103.8642
9.5161 | -103.8357
9.5198 | -103.8102
9.5235 | -103.7881
9.5273 | -103.7664
9.531 | -103.7525
9.5347 | -103.7536
9.5384 | -103.7612
9.5421 | -103.7737
9.5458 | -103.7938
9.5495 | -103.8207
9.5532 | -103.8579
9.5569 | -103.9067
9.5607 | -103.9656
9.5644 | -104.0303
9.5681 | -104.095
9.5718 | -104.1666
9.5755 | -104.2455
9.5792 | -104.3247
9.5829 | -104.3973
9.5866 | -104.4742
9.5904 | -104.5493
9.5941 | -104.6161
9.5978 | -104.6767
9.6015 | -104.7302
9.6052 | -104.7733
9.6089 | -104.812
9.6126 | -104.8434
9.6163 | -104.8628
9.62 | -104.8731
9.6238 | -104.8744
9.6275 | -104.8683
9.6312 | -104.8627
9.6349 | -104.8537
9.6386 | -104.8503
9.6423 | -104.8498
9.646 | -104.85
9.6497 | -104.8433
9.6534 | -104.841
9.6572 | -104.84
9.6609 | -104.8394
9.6646 | -104.8405
9.6683 | -104.8465
9.672 | -104.852
9.6757 | -104.8501
9.6794 | -104.8477
9.6831 | -104.8482
9.6868 | -104.8458
9.6906 | -104.8447
9.6943 | -104.8419
9.698 | -104.8368
9.7017 | -104.8305
9.7054 | -104.8206
9.7091 | -104.8159
9.7128 | -104.811
9.7165 | -104.801
9.7203 | -104.7896
9.724 | -104.782
9.7277 | -104.7715
9.7314 | -104.7593
9.7351 | -104.7486
9.7388 | -104.7373
9.7425 | -104.7272
9.7462 | -104.718
9.7499 | -104.7064
9.7537 | -104.6951
9.7574 | -104.6845
9.7611 | -104.6766
9.7648 | -104.6688
9.7685 | -104.6611
9.7722 | -104.6557
9.7759 | -104.6496
9.7796 | -104.6434
9.7833 | -104.6407
9.7871 | -104.636
9.7908 | -104.6316
9.7945 | -104.6344
9.7982 | -104.6394
9.8019 | -104.6447
9.8056 | -104.6511
9.8093 | -104.6572
9.813 | -104.663
9.8167 | -104.6671
9.8205 | -104.6798
9.8242 | -104.6937
9.8279 | -104.7063
9.8316 | -104.7143
9.8353 | -104.714
9.839 | -104.7191
9.8427 | -104.7244
9.8464 | -104.7263
9.8502 | -104.7299
9.8539 | -104.7224
9.8576 | -104.7124
9.8613 | -104.7049
9.865 | -104.694
9.8687 | -104.6746
9.8724 | -104.6547
9.8761 | -104.6343
9.8798 | -104.6132
9.8836 | -104.5883
9.8873 | -104.5632
9.891 | -104.5411
9.8947 | -104.5215
9.8984 | -104.5057
9.9021 | -104.4931
9.9058 | -104.4848
9.9095 | -104.4795
9.9132 | -104.4767
9.917 | -104.4782
9.9207 | -104.4826
9.9244 | -104.4934
9.9281 | -104.5055
9.9318 | -104.5207
9.9355 | -104.5364
9.9392 | -104.5491
9.9429 | -104.56
9.9466 | -104.5756
9.9504 | -104.5923
9.9541 | -104.6052
9.9578 | -104.6155
9.9615 | -104.6288
9.9652 | -104.6375
9.9689 | -104.6434
9.9726 | -104.6478
9.9763 | -104.6587
9.9801 | -104.6723
9.9838 | -104.6796
9.9875 | -104.6859
9.9912 | -104.699
9.9949 | -104.7136
9.9986 | -104.7335
10.0023 | -104.7499
10.006 | -104.7641
10.0097 | -104.7852
10.0135 | -104.8115
10.0172 | -104.8363
10.0209 | -104.8626
10.0246 | -104.8889
10.0283 | -104.919
10.032 | -104.9471
10.0357 | -104.9704
10.0394 | -104.9915
10.0431 | -105.0094
10.0469 | -105.0207
10.0506 | -105.0217
10.0543 | -105.0242
10.058 | -105.0284
10.0617 | -105.026
10.0654 | -105.0225
10.0691 | -105.0163
10.0728 | -105.0108
10.0765 | -105.0112
10.0803 | -105.0128
10.084 | -105.0096
10.0877 | -105.0094
10.0914 | -105.0075
10.0951 | -105.0099
10.0988 | -105.0164
10.1025 | -105.02
10.1062 | -105.0189
10.11 | -105.02
10.1137 | -105.0217
10.1174 | -105.0231
10.1211 | -105.0206
10.1248 | -105.019
10.1285 | -105.0149
10.1322 | -105.0117
10.1359 | -105.0053
10.1396 | -104.9993
10.1434 | -104.9946
10.1471 | -104.9922
10.1508 | -104.9871
10.1545 | -104.9778
10.1582 | -104.9715
10.1619 | -104.9644
10.1656 | -104.9567
10.1693 | -104.9508
10.173 | -104.945
10.1768 | -104.9372
10.1805 | -104.9308
10.1842 | -104.9268
10.1879 | -104.9197
10.1916 | -104.914
10.1953 | -104.9097
10.199 | -104.9084
10.2027 | -104.9066
10.2064 | -104.8992
10.2102 | -104.8929
10.2139 | -104.8883
10.2176 | -104.8837
10.2213 | -104.8739
10.225 | -104.8578
10.2287 | -104.8431
10.2324 | -104.8342
10.2361 | -104.8226
10.2399 | -104.81
10.2436 | -104.7943
10.2473 | -104.7785
10.251 | -104.7606
10.2547 | -104.7454
10.2584 | -104.7309
10.2621 | -104.7175
10.2658 | -104.7047
10.2695 | -104.6921
10.2733 | -104.6835
10.277 | -104.676
10.2807 | -104.6673
10.2844 | -104.6618
10.2881 | -104.6563
10.2918 | -104.6565
10.2955 | -104.6577
10.2992 | -104.6648
10.3029 | -104.6761
10.3067 | -104.6881
10.3104 | -104.7004
10.3141 | -104.7128
10.3178 | -104.7283
10.3215 | -104.7467
10.3252 | -104.7664
10.3289 | -104.7848
10.3326 | -104.8044
10.3363 | -104.8212
10.3401 | -104.8402
10.3438 | -104.8565
10.3475 | -104.8708
10.3512 | -104.8786
10.3549 | -104.8818
10.3586 | -104.8833
10.3623 | -104.8785
10.366 | -104.8699
10.3698 | -104.8552
10.3735 | -104.8426
10.3772 | -104.83
10.3809 | -104.8129
10.3846 | -104.7937
10.3883 | -104.7738
10.392 | -104.7549
10.3957 | -104.7362
10.3994 | -104.7161
10.4032 | -104.7036
10.4069 | -104.6866
10.4106 | -104.6655
10.4143 | -104.6522
10.418 | -104.6434
10.4217 | -104.6327
10.4254 | -104.6186
10.4291 | -104.6086
10.4328 | -104.6002
10.4366 | -104.593
10.4403 | -104.5864
10.444 | -104.5784
10.4477 | -104.5726
10.4514 | -104.564
10.4551 | -104.5547
10.4588 | -104.5467
10.4625 | -104.5402
10.4662 | -104.5362
10.47 | -104.5334
10.4737 | -104.5362
10.4774 | -104.5396
10.4811 | -104.5492
10.4848 | -104.5582
10.4885 | -104.5749
10.4922 | -104.5926
10.4959 | -104.6126
10.4997 | -104.6397
10.5034 | -104.6726
10.5071 | -104.7073
10.5108 | -104.7452
10.5145 | -104.7837
10.5182 | -104.8221
10.5219 | -104.8611
10.5256 | -104.8967
10.5293 | -104.932
10.5331 | -104.9672
10.5368 | -104.9992
10.5405 | -105.0301
10.5442 | -105.0578
10.5479 | -105.0874
10.5516 | -105.1145
10.5553 | -105.1366
10.559 | -105.1566
10.5627 | -105.1774
10.5665 | -105.1897
10.5702 | -105.1963
10.5739 | -105.2074
10.5776 | -105.2133
10.5813 | -105.211
10.585 | -105.207
10.5887 | -105.1927
10.5924 | -105.1741
10.5961 | -105.1517
10.5999 | -105.124
10.6036 | -105.0936
10.6073 | -105.0598
10.611 | -105.0272
10.6147 | -104.9981
10.6184 | -104.9668
10.6221 | -104.9303
10.6258 | -104.8934
10.6296 | -104.8603
10.6333 | -104.8273
10.637 | -104.8013
10.6407 | -104.7763
10.6444 | -104.7551
10.6481 | -104.7349
10.6518 | -104.7161
10.6555 | -104.7027
10.6592 | -104.6901
10.663 | -104.6795
10.6667 | -104.6716
10.6704 | -104.6668
10.6741 | -104.6668
10.6778 | -104.6653
10.6815 | -104.6619
10.6852 | -104.6599
10.6889 | -104.6593
10.6926 | -104.6597
10.6964 | -104.6579
10.7001 | -104.6542
10.7038 | -104.6531
10.7075 | -104.6534
10.7112 | -104.6549
10.7149 | -104.6565
10.7186 | -104.6576
10.7223 | -104.6584
10.7261 | -104.6625
10.7298 | -104.6685
10.7335 | -104.6828
10.7372 | -104.7005
10.7409 | -104.7219
10.7446 | -104.7483
10.7483 | -104.7749
10.752 | -104.8012
10.7557 | -104.8339
10.7595 | -104.8703
10.7632 | -104.9084
10.7669 | -104.9447
10.7706 | -104.9791
10.7743 | -105.0157
10.778 | -105.0568
10.7817 | -105.099
10.7854 | -105.1389
10.7891 | -105.1829
10.7929 | -105.2216
10.7966 | -105.2558
10.8003 | -105.2922
10.804 | -105.3287
10.8077 | -105.3609
10.8114 | -105.3907
10.8151 | -105.4188
10.8188 | -105.4469
10.8225 | -105.4692
10.8263 | -105.4848
10.83 | -105.4971
10.8337 | -105.5102
10.8374 | -105.517
10.8411 | -105.5133
10.8448 | -105.5011
10.8485 | -105.4861
10.8522 | -105.4718
10.8559 | -105.4484
10.8597 | -105.4201
10.8634 | -105.39
10.8671 | -105.3556
10.8708 | -105.3161
10.8745 | -105.2734
10.8782 | -105.2272
10.8819 | -105.1843
10.8856 | -105.1414
10.8894 | -105.0964
10.8931 | -105.0537
10.8968 | -105.0156
10.9005 | -104.9787
10.9042 | -104.9492
10.9079 | -104.9265
10.9116 | -104.9083
10.9153 | -104.8895
10.919 | -104.8733
10.9228 | -104.8619
10.9265 | -104.858
10.9302 | -104.8597
10.9339 | -104.8656
10.9376 | -104.8722
10.9413 | -104.8841
10.945 | -104.8967
10.9487 | -104.9081
10.9524 | -104.9249
10.9562 | -104.9442
10.9599 | -104.9623
10.9636 | -104.9801
10.9673 | -104.9983
10.971 | -105.0112
10.9747 | -105.0246
10.9784 | -105.0319
10.9821 | -105.0403
10.9858 | -105.0455
10.9896 | -105.0482
10.9933 | -105.0474
10.997 | -105.0468
11.0007 | -105.0507
11.0044 | -105.0565
11.0081 | -105.0573
11.0118 | -105.0531
11.0155 | -105.0542
11.0193 | -105.0554
11.023 | -105.0559
11.0267 | -105.0583
11.0304 | -105.0651
11.0341 | -105.077
11.0378 | -105.0884
11.0415 | -105.0992
11.0452 | -105.1137
11.0489 | -105.1311
11.0527 | -105.1477
11.0564 | -105.1643
11.0601 | -105.185
11.0638 | -105.2061
11.0675 | -105.2276
11.0712 | -105.2488
11.0749 | -105.2735
11.0786 | -105.2965
11.0823 | -105.3177
11.0861 | -105.3338
11.0898 | -105.3468
11.0935 | -105.3582
11.0972 | -105.3694
11.1009 | -105.3733
11.1046 | -105.3795
11.1083 | -105.3819
11.112 | -105.3764
11.1158 | -105.3644
11.1195 | -105.3423
11.1232 | -105.3154
11.1269 | -105.284
11.1306 | -105.2515
11.1343 | -105.2147
11.138 | -105.1779
11.1417 | -105.139
11.1454 | -105.0992
11.1492 | -105.0566
11.1529 | -105.0136
11.1566 | -104.9741
11.1603 | -104.9358
11.164 | -104.9047
11.1677 | -104.8792
11.1714 | -104.8592
11.1751 | -104.8453
11.1788 | -104.8354
11.1826 | -104.8284
11.1863 | -104.8266
11.19 | -104.8328
11.1937 | -104.8426
11.1974 | -104.858
11.2011 | -104.8733
11.2048 | -104.8865
11.2085 | -104.9037
11.2122 | -104.9189
11.216 | -104.9301
11.2197 | -104.9403
11.2234 | -104.9442
11.2271 | -104.9386
11.2308 | -104.9327
11.2345 | -104.9211
11.2382 | -104.9086
11.2419 | -104.8851
11.2457 | -104.8726
11.2494 | -104.8602
11.2531 | -104.8474
11.2568 | -104.8341
11.2605 | -104.8219
11.2642 | -104.8084
11.2679 | -104.7921
11.2716 | -104.779
11.2753 | -104.7736
11.2791 | -104.7702
11.2828 | -104.7692
11.2865 | -104.7695
11.2902 | -104.7723
11.2939 | -104.778
11.2976 | -104.7841
11.3013 | -104.7917
11.305 | -104.8028
11.3087 | -104.8189
11.3125 | -104.8361
11.3162 | -104.8544
11.3199 | -104.8733
11.3236 | -104.8941
11.3273 | -104.9133
11.331 | -104.9354
11.3347 | -104.9609
11.3384 | -104.9843
11.3421 | -105.0082
11.3459 | -105.036
11.3496 | -105.065
11.3533 | -105.0923
11.357 | -105.1192
11.3607 | -105.1478
11.3644 | -105.1753
11.3681 | -105.2056
11.3718 | -105.2344
11.3756 | -105.2651
11.3793 | -105.2962
11.383 | -105.3234
11.3867 | -105.3527
11.3904 | -105.3841
11.3941 | -105.4066
11.3978 | -105.4279
11.4015 | -105.4472
11.4052 | -105.4601
11.409 | -105.4703
11.4127 | -105.478
11.4164 | -105.4818
11.4201 | -105.4799
11.4238 | -105.4759
11.4275 | -105.4695
11.4312 | -105.4607
11.4349 | -105.4532
11.4386 | -105.4396
11.4424 | -105.4191
11.4461 | -105.3972
11.4498 | -105.3772
11.4535 | -105.3529
11.4572 | -105.3237
11.4609 | -105.295
11.4646 | -105.2701
11.4683 | -105.2469
11.472 | -105.2249
11.4758 | -105.2013
11.4795 | -105.1815
11.4832 | -105.1661
11.4869 | -105.1519
11.4906 | -105.1377
11.4943 | -105.1267
11.498 | -105.1187
11.5017 | -105.1125
11.5055 | -105.1113
11.5092 | -105.1126
11.5129 | -105.117
11.5166 | -105.1245
11.5203 | -105.1319
11.524 | -105.1433
11.5277 | -105.1563
11.5314 | -105.1731
11.5351 | -105.1897
11.5389 | -105.2097
11.5426 | -105.2311
11.5463 | -105.2534
11.55 | -105.2753
11.5537 | -105.2952
11.5574 | -105.3158
11.5611 | -105.3381
11.5648 | -105.356
11.5685 | -105.3702
11.5723 | -105.3859
11.576 | -105.4001
11.5797 | -105.4132
11.5834 | -105.4268
11.5871 | -105.4405
11.5908 | -105.4475
11.5945 | -105.4546
11.5982 | -105.4636
11.6019 | -105.4665
11.6057 | -105.4671
11.6094 | -105.4624
11.6131 | -105.4568
11.6168 | -105.449
11.6205 | -105.4437
11.6242 | -105.4399
11.6279 | -105.4307
11.6316 | -105.4174
11.6354 | -105.4012
11.6391 | -105.3828
11.6428 | -105.3675
11.6465 | -105.3534
11.6502 | -105.3376
11.6539 | -105.32
11.6576 | -105.3055
11.6613 | -105.2939
11.665 | -105.2792
11.6688 | -105.263
11.6725 | -105.2472
11.6762 | -105.2335
11.6799 | -105.2168
11.6836 | -105.2002
11.6873 | -105.1904
11.691 | -105.1806
11.6947 | -105.1734
11.6984 | -105.1684
11.7022 | -105.1642
11.7059 | -105.1581
11.7096 | -105.1539
11.7133 | -105.1526
11.717 | -105.1503
11.7207 | -105.1483
11.7244 | -105.146
11.7281 | -105.1435
11.7318 | -105.1397
11.7356 | -105.1389
11.7393 | -105.1405
11.743 | -105.1396
11.7467 | -105.1362
11.7504 | -105.1322
11.7541 | -105.1329
11.7578 | -105.1348
11.7615 | -105.1322
11.7653 | -105.1304
11.769 | -105.1285
11.7727 | -105.1265
11.7764 | -105.1172
11.7801 | -105.1157
11.7838 | -105.1097
11.7875 | -105.1011
11.7912 | -105.0906
11.7949 | -105.0779
11.7987 | -105.0646
11.8024 | -105.0521
11.8061 | -105.0402
11.8098 | -105.0288
11.8135 | -105.017
11.8172 | -105.0057
11.8209 | -104.9972
11.8246 | -104.9889
11.8283 | -104.9785
11.8321 | -104.9691
11.8358 | -104.9654
11.8395 | -104.9629
11.8432 | -104.9585
11.8469 | -104.9548
11.8506 | -104.9564
11.8543 | -104.9597
11.858 | -104.9599
11.8617 | -104.962
11.8655 | -104.9639
11.8692 | -104.9667
11.8729 | -104.9706
11.8766 | -104.9769
11.8803 | -104.9848
11.884 | -104.9954
11.8877 | -105.0063
11.8914 | -105.0169
11.8952 | -105.0276
11.8989 | -105.0379
11.9026 | -105.0517
11.9063 | -105.0634
11.91 | -105.0695
11.9137 | -105.0764
11.9174 | -105.0911
11.9211 | -105.1014
11.9248 | -105.1073
11.9286 | -105.1135
11.9323 | -105.1204
11.936 | -105.126
11.9397 | -105.1322
11.9434 | -105.1382
11.9471 | -105.1442
11.9508 | -105.1494
11.9545 | -105.1529
11.9582 | -105.1585
11.962 | -105.1609
11.9657 | -105.1621
11.9694 | -105.1635
11.9731 | -105.1637
11.9768 | -105.1598
11.9805 | -105.1535
11.9842 | -105.1473
11.9879 | -105.1428
11.9916 | -105.1369
11.9954 | -105.1315
11.9991 | -105.1292
12.0028 | -105.1297
12.0065 | -105.1297
12.0102 | -105.1309
12.0139 | -105.1339
12.0176 | -105.143
12.0213 | -105.1541
12.0251 | -105.1674
12.0288 | -105.1809
12.0325 | -105.1984
12.0362 | -105.2192
12.0399 | -105.2446
12.0436 | -105.2717
12.0473 | -105.3018
12.051 | -105.3326
12.0547 | -105.364
12.0585 | -105.3946
12.0622 | -105.426
12.0659 | -105.4574
12.0696 | -105.4881
12.0733 | -105.5168
12.077 | -105.543
12.0807 | -105.5505
12.0844 | -105.5482
12.0881 | -105.5549
12.0919 | -105.5508
12.0956 | -105.5471
12.0993 | -105.5367
12.103 | -105.5212
12.1067 | -105.5009
12.1104 | -105.4744
12.1141 | -105.4399
12.1178 | -105.399
12.1215 | -105.3559
12.1253 | -105.3094
12.129 | -105.2626
12.1327 | -105.2141
12.1364 | -105.1677
12.1401 | -105.1245
12.1438 | -105.0821
12.1475 | -105.0445
12.1512 | -105.0104
12.155 | -104.9835
12.1587 | -104.9632
12.1624 | -104.9478
12.1661 | -104.939
12.1698 | -104.9368
12.1735 | -104.9395
12.1772 | -104.9489
12.1809 | -104.9683
12.1846 | -104.9925
12.1884 | -105.0231
12.1921 | -105.0559
12.1958 | -105.0911
12.1995 | -105.1285
12.2032 | -105.1707
12.2069 | -105.2138
12.2106 | -105.254
12.2143 | -105.2922
12.218 | -105.3295
12.2218 | -105.3616
12.2255 | -105.3904
12.2292 | -105.417
12.2329 | -105.4415
12.2366 | -105.4609
12.2403 | -105.4782
12.244 | -105.4906
12.2477 | -105.499
12.2515 | -105.505
12.2552 | -105.5202
12.2589 | -105.5316
12.2626 | -105.5223
12.2663 | -105.5115
12.27 | -105.4994
12.2737 | -105.4903
12.2774 | -105.4791
12.2811 | -105.4665
12.2849 | -105.4567
12.2886 | -105.4496
12.2923 | -105.4452
12.296 | -105.4429
12.2997 | -105.4388
12.3034 | -105.4366
12.3071 | -105.4377
12.3108 | -105.4389
12.3145 | -105.4458
12.3183 | -105.4468
12.322 | -105.4506
12.3257 | -105.452
12.3294 | -105.4559
12.3331 | -105.4623
12.3368 | -105.4643
12.3405 | -105.4651
12.3442 | -105.4715
12.3479 | -105.4769
12.3517 | -105.4817
12.3554 | -105.4866
12.3591 | -105.4846
12.3628 | -105.4792
12.3665 | -105.4756
12.3702 | -105.4707
12.3739 | -105.4651
12.3776 | -105.4554
12.3813 | -105.444
12.3851 | -105.4321
12.3888 | -105.4198
12.3925 | -105.4031
12.3962 | -105.3847
12.3999 | -105.3665
12.4036 | -105.3493
12.4073 | -105.3295
12.411 | -105.3095
12.4148 | -105.2928
12.4185 | -105.2728
12.4222 | -105.2546
12.4259 | -105.2388
12.4296 | -105.2237
12.4333 | -105.2132
12.437 | -105.2061
12.4407 | -105.2004
12.4444 | -105.1961
12.4482 | -105.1946
12.4519 | -105.1948
12.4556 | -105.2003
12.4593 | -105.2095
12.463 | -105.2188
12.4667 | -105.23
12.4704 | -105.241
12.4741 | -105.2516
12.4778 | -105.2648
12.4816 | -105.281
12.4853 | -105.2984
12.489 | -105.319
12.4927 | -105.3416
12.4964 | -105.3639
12.5001 | -105.3837
12.5038 | -105.4035
12.5075 | -105.4261
12.5113 | -105.4472
12.515 | -105.4644
12.5187 | -105.4835
12.5224 | -105.5014
12.5261 | -105.5129
12.5298 | -105.5269
12.5335 | -105.5397
12.5372 | -105.5491
12.5409 | -105.5584
12.5447 | -105.5687
12.5484 | -105.5738
12.5521 | -105.5779
12.5558 | -105.5794
12.5595 | -105.5802
12.5632 | -105.5782
12.5669 | -105.573
12.5706 | -105.5654
12.5743 | -105.5615
12.5781 | -105.5585
12.5818 | -105.5537
12.5855 | -105.545
12.5892 | -105.5388
12.5929 | -105.5329
12.5966 | -105.5222
12.6003 | -105.517
12.604 | -105.5163
12.6077 | -105.5121
12.6115 | -105.5046
12.6152 | -105.4957
12.6189 | -105.4909
12.6226 | -105.4857
12.6263 | -105.4808
12.63 | -105.4785
12.6337 | -105.4751
12.6374 | -105.4705
12.6412 | -105.4667
12.6449 | -105.4636
12.6486 | -105.4616
12.6523 | -105.4559
12.656 | -105.449
12.6597 | -105.4421
12.6634 | -105.4328
12.6671 | -105.4238
12.6708 | -105.4145
12.6746 | -105.4048
12.6783 | -105.3958
12.682 | -105.3878
12.6857 | -105.3768
12.6894 | -105.3661
12.6931 | -105.3571
12.6968 | -105.3454
12.7005 | -105.3362
12.7042 | -105.3289
12.708 | -105.3248
12.7117 | -105.3178
12.7154 | -105.3137
12.7191 | -105.3108
12.7228 | -105.3084
12.7265 | -105.3083
12.7302 | -105.3063
12.7339 | -105.304
12.7376 | -105.3039
12.7414 | -105.3048
12.7451 | -105.3065
12.7488 | -105.3084
12.7525 | -105.3098
12.7562 | -105.3087
12.7599 | -105.3054
12.7636 | -105.3014
12.7673 | -105.2965
12.7711 | -105.2897
12.7748 | -105.2841
12.7785 | -105.2777
12.7822 | -105.261
12.7859 | -105.2452
12.7896 | -105.2337
12.7933 | -105.2238
12.797 | -105.2122
12.8007 | -105.1975
12.8045 | -105.1823
12.8082 | -105.1676
12.8119 | -105.1515
12.8156 | -105.1366
12.8193 | -105.1203
12.823 | -105.1064
12.8267 | -105.0935
12.8304 | -105.0845
12.8341 | -105.0784
12.8379 | -105.0716
12.8416 | -105.0674
12.8453 | -105.0656
12.849 | -105.0775
12.8527 | -105.0789
12.8564 | -105.082
12.8601 | -105.0882
12.8638 | -105.099
12.8675 | -105.1123
12.8713 | -105.1298
12.875 | -105.1484
12.8787 | -105.164
12.8824 | -105.1842
12.8861 | -105.2034
12.8898 | -105.2225
12.8935 | -105.2426
12.8972 | -105.2637
12.901 | -105.2842
12.9047 | -105.303
12.9084 | -105.3209
12.9121 | -105.3375
12.9158 | -105.3531
12.9195 | -105.3608
12.9232 | -105.3655
12.9269 | -105.3706
12.9306 | -105.374
12.9344 | -105.3749
12.9381 | -105.3726
12.9418 | -105.3666
12.9455 | -105.3596
12.9492 | -105.3548
12.9529 | -105.3488
12.9566 | -105.343
12.9603 | -105.3364
12.964 | -105.3334
12.9678 | -105.3352
12.9715 | -105.333
12.9752 | -105.3319
12.9789 | -105.3326
12.9826 | -105.3351
12.9863 | -105.3406
12.99 | -105.3492
12.9937 | -105.3556
12.9974 | -105.3645
13.0012 | -105.3775
13.0049 | -105.3956
13.0086 | -105.4151
13.0123 | -105.4335
13.016 | -105.4492
13.0197 | -105.4698
13.0234 | -105.4939
13.0271 | -105.5163
13.0309 | -105.5397
13.0346 | -105.5637
13.0383 | -105.586
13.042 | -105.6078
13.0457 | -105.6293
13.0494 | -105.6484
13.0531 | -105.6667
13.0568 | -105.6896
13.0605 | -105.7065
13.0643 | -105.724
13.068 | -105.7435
13.0717 | -105.7597
13.0754 | -105.7746
13.0791 | -105.7849
13.0828 | -105.7927
13.0865 | -105.7966
13.0902 | -105.7994
13.0939 | -105.8031
13.0977 | -105.8046
13.1014 | -105.8014
13.1051 | -105.7988
13.1088 | -105.8007
13.1125 | -105.7955
13.1162 | -105.7831
13.1199 | -105.7697
13.1236 | -105.7586
13.1273 | -105.7481
13.1311 | -105.7352
13.1348 | -105.7196
13.1385 | -105.7003
13.1422 | -105.6834
13.1459 | -105.6629
13.1496 | -105.6378
13.1533 | -105.6159
13.157 | -105.5945
13.1608 | -105.5732
13.1645 | -105.5528
13.1682 | -105.5333
13.1719 | -105.5157
13.1756 | -105.4966
13.1793 | -105.4802
13.183 | -105.468
13.1867 | -105.4533
13.1904 | -105.4417
13.1942 | -105.4307
13.1979 | -105.4193
13.2016 | -105.4132
13.2053 | -105.4106
13.209 | -105.4072
13.2127 | -105.404
13.2164 | -105.4039
13.2201 | -105.4063
13.2238 | -105.413
13.2276 | -105.4246
13.2313 | -105.4375
13.235 | -105.4526
13.2387 | -105.4698
13.2424 | -105.487
13.2461 | -105.5022
13.2498 | -105.5226
13.2535 | -105.5399
13.2572 | -105.5579
13.261 | -105.5752
13.2647 | -105.5937
13.2684 | -105.6146
13.2721 | -105.6324
13.2758 | -105.6521
13.2795 | -105.6677
13.2832 | -105.6777
13.2869 | -105.6879
13.2907 | -105.6979
13.2944 | -105.7056
13.2981 | -105.7088
13.3018 | -105.7152
13.3055 | -105.7202
13.3092 | -105.7188
13.3129 | -105.715
13.3166 | -105.7095
13.3203 | -105.7002
13.3241 | -105.6924
13.3278 | -105.6857
13.3315 | -105.6761
13.3352 | -105.6699
13.3389 | -105.6662
13.3426 | -105.6579
13.3463 | -105.6483
13.35 | -105.6365
13.3537 | -105.6226
13.3575 | -105.6106
13.3612 | -105.5965
13.3649 | -105.5839
13.3686 | -105.5682
13.3723 | -105.5515
13.376 | -105.5355
13.3797 | -105.5171
13.3834 | -105.4998
13.3871 | -105.4783
13.3909 | -105.4593
13.3946 | -105.4386
13.3983 | -105.4192
13.402 | -105.4009
13.4057 | -105.3851
13.4094 | -105.3668
13.4131 | -105.3513
13.4168 | -105.3337
13.4206 | -105.3136
13.4243 | -105.2973
13.428 | -105.2838
13.4317 | -105.2711
13.4354 | -105.259
13.4391 | -105.2484
13.4428 | -105.2408
13.4465 | -105.2356
13.4502 | -105.2307
13.454 | -105.2278
13.4577 | -105.2254
13.4614 | -105.2226
13.4651 | -105.2222
13.4688 | -105.223
13.4725 | -105.2259
13.4762 | -105.2347
13.4799 | -105.2466
13.4836 | -105.259
13.4874 | -105.2731
13.4911 | -105.2909
13.4948 | -105.3094
13.4985 | -105.3282
13.5022 | -105.3533
13.5059 | -105.379
13.5096 | -105.4063
13.5133 | -105.4355
13.517 | -105.4625
13.5208 | -105.4882
13.5245 | -105.5131
13.5282 | -105.5346
13.5319 | -105.5536
13.5356 | -105.5727
13.5393 | -105.5905
13.543 | -105.6082
13.5467 | -105.6232
13.5505 | -105.6358
13.5542 | -105.6468
13.5579 | -105.6524
13.5616 | -105.6562
13.5653 | -105.659
13.569 | -105.6601
13.5727 | -105.6581
13.5764 | -105.6544
13.5801 | -105.6503
13.5839 | -105.6475
13.5876 | -105.6412
13.5913 | -105.6343
13.595 | -105.6234
13.5987 | -105.6134
13.6024 | -105.5978
13.6061 | -105.5858
13.6098 | -105.5715
13.6135 | -105.5593
13.6173 | -105.5499
13.621 | -105.5389
13.6247 | -105.5272
13.6284 | -105.5147
13.6321 | -105.5029
13.6358 | -105.4899
13.6395 | -105.476
13.6432 | -105.4614
13.647 | -105.4494
13.6507 | -105.4385
13.6544 | -105.4279
13.6581 | -105.4204
13.6618 | -105.4123
13.6655 | -105.402
13.6692 | -105.3951
13.6729 | -105.3874
13.6766 | -105.3802
13.6804 | -105.3729
13.6841 | -105.3677
13.6878 | -105.3654
13.6915 | -105.3623
13.6952 | -105.3589
13.6989 | -105.3549
13.7026 | -105.349
13.7063 | -105.3451
13.71 | -105.3443
13.7138 | -105.3452
13.7175 | -105.3464
13.7212 | -105.3479
13.7249 | -105.353
13.7286 | -105.3595
13.7323 | -105.3656
13.736 | -105.3717
13.7397 | -105.3776
13.7434 | -105.3852
13.7472 | -105.3906
13.7509 | -105.3962
13.7546 | -105.4042
13.7583 | -105.4124
13.762 | -105.4218
13.7657 | -105.4327
13.7694 | -105.4433
13.7731 | -105.4525
13.7768 | -105.4646
13.7806 | -105.4778
13.7843 | -105.4874
13.788 | -105.4993
13.7917 | -105.5102
13.7954 | -105.5213
13.7991 | -105.5359
13.8028 | -105.5457
13.8065 | -105.555
13.8103 | -105.5643
13.814 | -105.5694
13.8177 | -105.5697
13.8214 | -105.5697
13.8251 | -105.5738
13.8288 | -105.5755
13.8325 | -105.5778
13.8362 | -105.5749
13.8399 | -105.5723
13.8437 | -105.5681
13.8474 | -105.5594
13.8511 | -105.5509
13.8548 | -105.5395
13.8585 | -105.5284
13.8622 | -105.5172
13.8659 | -105.5023
13.8696 | -105.4872
13.8733 | -105.4721
13.8771 | -105.4564
13.8808 | -105.4408
13.8845 | -105.4251
13.8882 | -105.4109
13.8919 | -105.3999
13.8956 | -105.3838
13.8993 | -105.3694
13.903 | -105.3531
13.9068 | -105.3356
13.9105 | -105.3183
13.9142 | -105.3011
13.9179 | -105.2802
13.9216 | -105.259
13.9253 | -105.2372
13.929 | -105.2163
13.9327 | -105.1965
13.9364 | -105.177
13.9402 | -105.1547
13.9439 | -105.1341
13.9476 | -105.1164
13.9513 | -105.0944
13.955 | -105.0672
13.9587 | -105.0287
13.9624 | -104.9674
13.9661 | -104.9219
13.9698 | -104.8879
13.9736 | -104.8506
13.9773 | -104.8115
13.981 | -104.7733
13.9847 | -104.7359
13.9884 | -104.6948
13.9921 | -104.6554
13.9958 | -104.6121
13.9995 | -104.5666
14.0032 | -104.5235
14.007 | -104.4792
14.0107 | -104.4345
14.0144 | -104.3935
14.0181 | -104.3552
14.0218 | -104.3156
14.0255 | -104.2767
14.0292 | -104.241
14.0329 | -104.2077
14.0366 | -104.1826
14.0404 | -104.1603
14.0441 | -104.1426
14.0478 | -104.1317
14.0515 | -104.1262
14.0552 | -104.1288
14.0589 | -104.1359
14.0626 | -104.1527
14.0663 | -104.1758
14.0701 | -104.2059
14.0738 | -104.2397
14.0775 | -104.2799
14.0812 | -104.3245
14.0849 | -104.3733
14.0886 | -104.4272
14.0923 | -104.4835
14.096 | -104.5421
14.0997 | -104.598
14.1035 | -104.6557
14.1072 | -104.7159
14.1109 | -104.774
14.1146 | -104.8304
14.1183 | -104.8837
14.122 | -104.9345
14.1257 | -104.9854
14.1294 | -105.0287
14.1331 | -105.0665
14.1369 | -105.0991
14.1406 | -105.1293
14.1443 | -105.1583
14.148 | -105.1837
14.1517 | -105.2052
14.1554 | -105.2291
14.1591 | -105.2525
14.1628 | -105.2805
14.1666 | -105.3185
14.1703 | -105.3292
14.174 | -105.3274
14.1777 | -105.3284
14.1814 | -105.33
14.1851 | -105.3304
14.1888 | -105.3304
14.1925 | -105.3329
14.1962 | -105.3357
14.2 | -105.3373
14.2037 | -105.3403
14.2074 | -105.3426
14.2111 | -105.3467
14.2148 | -105.3507
14.2185 | -105.3523
14.2222 | -105.3551
14.2259 | -105.3583
14.2296 | -105.357
14.2334 | -105.356
14.2371 | -105.3548
14.2408 | -105.3505
14.2445 | -105.3452
14.2482 | -105.3354
14.2519 | -105.325
14.2556 | -105.3163
14.2593 | -105.3065
14.263 | -105.2961
14.2668 | -105.2804
14.2705 | -105.2636
14.2742 | -105.2484
14.2779 | -105.2387
14.2816 | -105.229
14.2853 | -105.2153
14.289 | -105.2033
14.2927 | -105.1911
14.2965 | -105.1828
14.3002 | -105.1791
14.3039 | -105.1785
14.3076 | -105.1768
14.3113 | -105.1752
14.315 | -105.1728
14.3187 | -105.1698
14.3224 | -105.1709
14.3261 | -105.1717
14.3299 | -105.1726
14.3336 | -105.1752
14.3373 | -105.1804
14.341 | -105.1879
14.3447 | -105.1924
14.3484 | -105.1971
14.3521 | -105.2015
14.3558 | -105.2053
14.3595 | -105.2107
14.3633 | -105.221
14.367 | -105.2315
14.3707 | -105.2416
14.3744 | -105.2505
14.3781 | -105.2565
14.3818 | -105.2643
14.3855 | -105.2712
14.3892 | -105.2799
14.3929 | -105.2902
14.3967 | -105.2988
14.4004 | -105.3029
14.4041 | -105.3107
14.4078 | -105.3165
14.4115 | -105.3189
14.4152 | -105.3218
14.4189 | -105.3256
14.4226 | -105.3284
14.4264 | -105.3277
14.4301 | -105.3283
14.4338 | -105.3326
14.4375 | -105.3327
14.4412 | -105.3321
14.4449 | -105.3302
14.4486 | -105.327
14.4523 | -105.3266
14.456 | -105.3203
14.4598 | -105.3188
14.4635 | -105.3153
14.4672 | -105.3107
14.4709 | -105.3078
14.4746 | -105.3076
14.4783 | -105.3093
14.482 | -105.3053
14.4857 | -105.2996
14.4894 | -105.2963
14.4932 | -105.2979
14.4969 | -105.2979
14.5006 | -105.2995
14.5043 | -105.2983
14.508 | -105.2976
14.5117 | -105.2986
14.5154 | -105.3021
14.5191 | -105.3052
14.5228 | -105.3123
14.5266 | -105.3221
14.5303 | -105.3316
14.534 | -105.3441
14.5377 | -105.3555
14.5414 | -105.3684
14.5451 | -105.3807
14.5488 | -105.3907
14.5525 | -105.4022
14.5563 | -105.4154
14.56 | -105.43
14.5637 | -105.4454
14.5674 | -105.4605
14.5711 | -105.47
14.5748 | -105.4824
14.5785 | -105.4987
14.5822 | -105.5175
14.5859 | -105.5328
14.5897 | -105.5481
14.5934 | -105.5641
14.5971 | -105.577
14.6008 | -105.59
14.6045 | -105.6032
14.6082 | -105.6157
14.6119 | -105.6259
14.6156 | -105.6354
14.6193 | -105.6437
14.6231 | -105.6489
14.6268 | -105.6535
14.6305 | -105.6575
14.6342 | -105.6584
14.6379 | -105.6551
14.6416 | -105.6481
14.6453 | -105.639
14.649 | -105.6283
14.6527 | -105.6203
14.6565 | -105.6119
14.6602 | -105.6018
14.6639 | -105.5927
14.6676 | -105.5816
14.6713 | -105.5681
14.675 | -105.5528
14.6787 | -105.5372
14.6824 | -105.5246
14.6862 | -105.5137
14.6899 | -105.5016
14.6936 | -105.4905
14.6973 | -105.4798
14.701 | -105.4692
14.7047 | -105.4581
14.7084 | -105.4466
14.7121 | -105.4366
14.7158 | -105.4303
14.7196 | -105.4254
14.7233 | -105.422
14.727 | -105.4146
14.7307 | -105.4086
14.7344 | -105.4055
14.7381 | -105.4029
14.7418 | -105.4014
14.7455 | -105.4014
14.7492 | -105.4019
14.753 | -105.4024
14.7567 | -105.4016
14.7604 | -105.4035
14.7641 | -105.4075
14.7678 | -105.4096
14.7715 | -105.4104
14.7752 | -105.4134
14.7789 | -105.4169
14.7826 | -105.4239
14.7864 | -105.4281
14.7901 | -105.4324
14.7938 | -105.435
14.7975 | -105.4421
14.8012 | -105.4467
14.8049 | -105.4492
14.8086 | -105.4537
14.8123 | -105.4577
14.8161 | -105.4587
14.8198 | -105.4644
14.8235 | -105.4676
14.8272 | -105.4712
14.8309 | -105.4757
14.8346 | -105.4809
14.8383 | -105.4846
14.842 | -105.4898
14.8457 | -105.4988
14.8495 | -105.51
14.8532 | -105.5203
14.8569 | -105.5308
14.8606 | -105.5392
14.8643 | -105.542
14.868 | -105.5454
14.8717 | -105.5536
14.8754 | -105.5634
14.8791 | -105.5772
14.8829 | -105.5882
14.8866 | -105.5999
14.8903 | -105.6105
14.894 | -105.6159
14.8977 | -105.6224
14.9014 | -105.6318
14.9051 | -105.6417
14.9088 | -105.6474
14.9125 | -105.6514
14.9163 | -105.6539
14.92 | -105.6563
14.9237 | -105.6543
14.9274 | -105.6486
14.9311 | -105.6396
14.9348 | -105.6302
14.9385 | -105.6221
14.9422 | -105.6127
14.946 | -105.5985
14.9497 | -105.5825
14.9534 | -105.566
14.9571 | -105.5498
14.9608 | -105.5361
14.9645 | -105.5242
14.9682 | -105.5127
14.9719 | -105.5002
14.9756 | -105.4893
14.9794 | -105.4812
14.9831 | -105.4739
14.9868 | -105.4647
14.9905 | -105.4103
14.9942 | -105.3957
14.9979 | -105.3814
15.0016 | -105.3672
15.0053 | -105.3496
15.009 | -105.3327
15.0128 | -105.3162
15.0165 | -105.2973
15.0202 | -105.2781
15.0239 | -105.2576
15.0276 | -105.2358
15.0313 | -105.2136
15.035 | -105.1865
15.0387 | -105.1579
15.0425 | -105.1261
15.0462 | -105.0933
15.0499 | -105.0587
15.0536 | -105.0251
15.0573 | -104.9911
15.061 | -104.9588
15.0647 | -104.9258
15.0684 | -104.8965
15.0721 | -104.8706
15.0759 | -104.8475
15.0796 | -104.8276
15.0833 | -104.8112
15.087 | -104.7929
15.0907 | -104.7767
15.0944 | -104.7635
15.0981 | -104.7592
15.1018 | -104.759
15.1055 | -104.7633
15.1093 | -104.7741
15.113 | -104.7835
15.1167 | -104.7934
15.1204 | -104.8077
15.1241 | -104.8269
15.1278 | -104.8458
15.1315 | -104.8658
15.1352 | -104.8871
15.1389 | -104.9099
15.1427 | -104.9335
15.1464 | -104.9577
15.1501 | -104.9801
15.1538 | -105.0027
15.1575 | -105.0245
15.1612 | -105.0476
15.1649 | -105.0688
15.1686 | -105.0887
15.1723 | -105.1068
15.1761 | -105.1205
15.1798 | -105.1348
15.1835 | -105.1497
15.1872 | -105.1642
15.1909 | -105.1767
15.1946 | -105.188
15.1983 | -105.2002
15.202 | -105.2124
15.2058 | -105.2466
15.2095 | -105.2685
15.2132 | -105.2695
15.2169 | -105.2719
15.2206 | -105.2751
15.2243 | -105.2808
15.228 | -105.2867
15.2317 | -105.2886
15.2354 | -105.2919
15.2392 | -105.2977
15.2429 | -105.3006
15.2466 | -105.3019
15.2503 | -105.3053
15.254 | -105.3114
15.2577 | -105.3157
15.2614 | -105.3206
15.2651 | -105.3262
15.2688 | -105.3323
15.2726 | -105.3333
15.2763 | -105.3347
15.28 | -105.334
15.2837 | -105.3358
15.2874 | -105.3382
15.2911 | -105.3402
15.2948 | -105.3411
15.2985 | -105.3425
15.3023 | -105.3433
15.306 | -105.3486
15.3097 | -105.3522
15.3134 | -105.3585
15.3171 | -105.3623
15.3208 | -105.3652
15.3245 | -105.3724
15.3282 | -105.3803
15.3319 | -105.3894
15.3357 | -105.401
15.3394 | -105.4153
15.3431 | -105.4302
15.3468 | -105.4469
15.3505 | -105.4616
15.3542 | -105.4767
15.3579 | -105.4924
15.3616 | -105.5091
15.3653 | -105.5278
15.3691 | -105.5454
15.3728 | -105.5631
15.3765 | -105.5797
15.3802 | -105.5936
15.3839 | -105.6089
15.3876 | -105.6191
15.3913 | -105.6279
15.395 | -105.6403
15.3987 | -105.6495
15.4025 | -105.655
15.4062 | -105.6579
15.4099 | -105.6595
15.4136 | -105.6586
15.4173 | -105.6537
15.421 | -105.6471
15.4247 | -105.6369
15.4284 | -105.6255
15.4322 | -105.6121
15.4359 | -105.5973
15.4396 | -105.5835
15.4433 | -105.5671
15.447 | -105.5516
15.4507 | -105.5363
15.4544 | -105.5178
15.4581 | -105.4962
15.4618 | -105.4791
15.4656 | -105.4614
15.4693 | -105.4388
15.473 | -105.4174
15.4767 | -105.3983
15.4804 | -105.3778
15.4841 | -105.3574
15.4878 | -105.3364
15.4915 | -105.3189
15.4952 | -105.3016
15.499 | -105.2858
15.5027 | -105.2695
15.5064 | -105.253
15.5101 | -105.2412
15.5138 | -105.2315
15.5175 | -105.2215
15.5212 | -105.2121
15.5249 | -105.2098
15.5286 | -105.2006
15.5324 | -105.1976
15.5361 | -105.1949
15.5398 | -105.1947
15.5435 | -105.1949
15.5472 | -105.1945
15.5509 | -105.1957
15.5546 | -105.1981
15.5583 | -105.2021
15.5621 | -105.2063
15.5658 | -105.2091
15.5695 | -105.2136
15.5732 | -105.2215
15.5769 | -105.2283
15.5806 | -105.2347
15.5843 | -105.2409
15.588 | -105.2457
15.5917 | -105.2509
15.5955 | -105.2556
15.5992 | -105.2612
15.6029 | -105.2667
15.6066 | -105.2728
15.6103 | -105.2799
15.614 | -105.2846
15.6177 | -105.2865
15.6214 | -105.2881
15.6251 | -105.2912
15.6289 | -105.2948
15.6326 | -105.2969
15.6363 | -105.299
15.64 | -105.3021
15.6437 | -105.3066
15.6474 | -105.3108
15.6511 | -105.3143
15.6548 | -105.3185
15.6585 | -105.3221
15.6623 | -105.324
15.666 | -105.3244
15.6697 | -105.3241
15.6734 | -105.3267
15.6771 | -105.3329
15.6808 | -105.3371
15.6845 | -105.3358
15.6882 | -105.3391
15.692 | -105.3443
15.6957 | -105.3465
15.6994 | -105.3502
15.7031 | -105.3541
15.7068 | -105.3581
15.7105 | -105.3632
15.7142 | -105.3699
15.7179 | -105.3768
15.7216 | -105.3849
15.7254 | -105.3934
15.7291 | -105.4019
15.7328 | -105.4078
15.7365 | -105.4139
15.7402 | -105.4221
15.7439 | -105.431
15.7476 | -105.4395
15.7513 | -105.4505
15.755 | -105.4571
15.7588 | -105.4564
15.7625 | -105.4605
15.7662 | -105.4652
15.7699 | -105.4719
15.7736 | -105.4754
15.7773 | -105.4753
15.781 | -105.4757
15.7847 | -105.4726
15.7884 | -105.4673
15.7922 | -105.4607
15.7959 | -105.4549
15.7996 | -105.4444
15.8033 | -105.4308
15.807 | -105.4169
15.8107 | -105.3976
15.8144 | -105.3764
15.8181 | -105.3565
15.8219 | -105.3355
15.8256 | -105.3117
15.8293 | -105.2877
15.833 | -105.2609
15.8367 | -105.236
15.8404 | -105.2114
15.8441 | -105.1851
15.8478 | -105.1591
15.8515 | -105.1317
15.8553 | -105.1041
15.859 | -105.0804
15.8627 | -105.0575
15.8664 | -105.0359
15.8701 | -105.0157
15.8738 | -104.9996
15.8775 | -104.9833
15.8812 | -104.9671
15.8849 | -104.9532
15.8887 | -104.9439
15.8924 | -104.9379
15.8961 | -104.9322
15.8998 | -104.9272
15.9035 | -104.9235
15.9072 | -104.9247
15.9109 | -104.9297
15.9146 | -104.9328
15.9183 | -104.9369
15.9221 | -104.9452
15.9258 | -104.9516
15.9295 | -104.9591
15.9332 | -104.9692
15.9369 | -104.9799
15.9406 | -104.9901
15.9443 | -105.0001
15.948 | -105.0121
15.9518 | -105.0248
15.9555 | -105.0391
15.9592 | -105.0536
15.9629 | -105.0688
15.9666 | -105.0833
15.9703 | -105.0967
15.974 | -105.1105
15.9777 | -105.1244
15.9814 | -105.1403
15.9852 | -105.1609
15.9889 | -105.1778
15.9926 | -105.1911
15.9963 | -105.2042
16 | -105.2169
16.0037 | -105.2295
16.0074 | -105.2423
16.0111 | -105.2558
16.0148 | -105.2683
16.0186 | -105.2835
16.0223 | -105.2956
16.026 | -105.3066
16.0297 | -105.3168
16.0334 | -105.3266
16.0371 | -105.3416
16.0408 | -105.3555
16.0445 | -105.3677
16.0482 | -105.3796
16.052 | -105.3907
16.0557 | -105.4008
16.0594 | -105.4109
16.0631 | -105.4187
16.0668 | -105.4246
16.0705 | -105.4309
16.0742 | -105.4382
16.0779 | -105.4429
16.0817 | -105.4477
16.0854 | -105.4537
16.0891 | -105.4552
16.0928 | -105.4562
16.0965 | -105.4566
16.1002 | -105.4548
16.1039 | -105.4517
16.1076 | -105.4499
16.1113 | -105.4494
16.1151 | -105.4451
16.1188 | -105.4369
16.1225 | -105.4307
16.1262 | -105.4252
16.1299 | -105.4206
16.1336 | -105.4161
16.1373 | -105.4086
16.141 | -105.4012
16.1447 | -105.3935
16.1485 | -105.387
16.1522 | -105.3782
16.1559 | -105.3708
16.1596 | -105.366
16.1633 | -105.3599
16.167 | -105.3551
16.1707 | -105.3515
16.1744 | -105.3489
16.1781 | -105.3477
16.1819 | -105.3466
16.1856 | -105.3421
16.1893 | -105.3419
16.193 | -105.3445
16.1967 | -105.3472
16.2004 | -105.3493
16.2041 | -105.3518
16.2078 | -105.3534
16.2116 | -105.3575
16.2153 | -105.3586
16.219 | -105.3588
16.2227 | -105.3651
16.2264 | -105.3709
16.2301 | -105.3785
16.2338 | -105.3861
16.2375 | -105.3903
16.2412 | -105.3924
16.245 | -105.3935
16.2487 | -105.3925
16.2524 | -105.3913
16.2561 | -105.3911
16.2598 | -105.3894
16.2635 | -105.3859
16.2672 | -105.3796
16.2709 | -105.3717
16.2746 | -105.3615
16.2784 | -105.3498
16.2821 | -105.3416
16.2858 | -105.3334
16.2895 | -105.3233
16.2932 | -105.313
16.2969 | -105.303
16.3006 | -105.2925
16.3043 | -105.2827
16.308 | -105.2716
16.3118 | -105.2643
16.3155 | -105.2572
16.3192 | -105.2456
16.3229 | -105.2366
16.3266 | -105.2292
16.3303 | -105.2209
16.334 | -105.2119
16.3377 | -105.2044
16.3415 | -105.1974
16.3452 | -105.1943
16.3489 | -105.1944
16.3526 | -105.1964
16.3563 | -105.1972
16.36 | -105.1977
16.3637 | -105.1984
16.3674 | -105.1986
16.3711 | -105.1995
16.3749 | -105.2027
16.3786 | -105.2092
16.3823 | -105.2158
16.386 | -105.2209
16.3897 | -105.227
16.3934 | -105.2339
16.3971 | -105.2413
16.4008 | -105.2475
16.4045 | -105.2565
16.4083 | -105.2644
16.412 | -105.2704
16.4157 | -105.2776
16.4194 | -105.29
16.4231 | -105.2985
16.4268 | -105.3035
16.4305 | -105.3097
16.4342 | -105.3186
16.4379 | -105.3243
16.4417 | -105.3283
16.4454 | -105.3307
16.4491 | -105.3347
16.4528 | -105.3386
16.4565 | -105.3364
16.4602 | -105.3342
16.4639 | -105.3315
16.4676 | -105.3302
16.4714 | -105.3303
16.4751 | -105.331
16.4788 | -105.3327
16.4825 | -105.3339
16.4862 | -105.3334
16.4899 | -105.3343
16.4936 | -105.3358
16.4973 | -105.3365
16.501 | -105.3367
16.5048 | -105.3374
16.5085 | -105.3378
16.5122 | -105.3398
16.5159 | -105.3397
16.5196 | -105.341
16.5233 | -105.3456
16.527 | -105.3482
16.5307 | -105.3556
16.5344 | -105.3637
16.5382 | -105.3701
16.5419 | -105.38
16.5456 | -105.3886
16.5493 | -105.3961
16.553 | -105.4064
16.5567 | -105.4195
16.5604 | -105.4319
16.5641 | -105.4461
16.5678 | -105.4627
16.5716 | -105.4777
16.5753 | -105.4898
16.579 | -105.4996
16.5827 | -105.5111
16.5864 | -105.5241
16.5901 | -105.5384
16.5938 | -105.5534
16.5975 | -105.5672
16.6013 | -105.582
16.605 | -105.5965
16.6087 | -105.6123
16.6124 | -105.6253
16.6161 | -105.633
16.6198 | -105.641
16.6235 | -105.6508
16.6272 | -105.6618
16.6309 | -105.6676
16.6347 | -105.6726
16.6384 | -105.6767
16.6421 | -105.6796
16.6458 | -105.6824
16.6495 | -105.6839
16.6532 | -105.6828
16.6569 | -105.6812
16.6606 | -105.6785
16.6643 | -105.6738
16.6681 | -105.6707
16.6718 | -105.6657
16.6755 | -105.6627
16.6792 | -105.6596
16.6829 | -105.6541
16.6866 | -105.6492
16.6903 | -105.6426
16.694 | -105.6384
16.6977 | -105.6327
16.7015 | -105.6244
16.7052 | -105.6132
16.7089 | -105.6025
16.7126 | -105.594
16.7163 | -105.5848
16.72 | -105.5753
16.7237 | -105.5708
16.7274 | -105.5662
16.7312 | -105.5602
16.7349 | -105.5556
16.7386 | -105.5506
16.7423 | -105.5464
16.746 | -105.5461
16.7497 | -105.5446
16.7534 | -105.5444
16.7571 | -105.5438
16.7608 | -105.5419
16.7646 | -105.5406
16.7683 | -105.5376
16.772 | -105.5386
16.7757 | -105.5408
16.7794 | -105.542
16.7831 | -105.5441
16.7868 | -105.5499
16.7905 | -105.5572
16.7942 | -105.5602
16.798 | -105.5628
16.8017 | -105.5667
16.8054 | -105.5739
16.8091 | -105.5814
16.8128 | -105.5922
16.8165 | -105.6005
16.8202 | -105.6062
16.8239 | -105.612
16.8276 | -105.6191
16.8314 | -105.6264
16.8351 | -105.6317
16.8388 | -105.6363
16.8425 | -105.6411
16.8462 | -105.6465
16.8499 | -105.6489
16.8536 | -105.6495
16.8573 | -105.6548
16.8611 | -105.6565
16.8648 | -105.6582
16.8685 | -105.657
16.8722 | -105.6564
16.8759 | -105.6548
16.8796 | -105.6546
16.8833 | -105.6505
16.887 | -105.6482
16.8907 | -105.6461
16.8945 | -105.6433
16.8982 | -105.6392
16.9019 | -105.6365
16.9056 | -105.6356
16.9093 | -105.6327
16.913 | -105.6293
16.9167 | -105.627
16.9204 | -105.6265
16.9241 | -105.6257
16.9279 | -105.6236
16.9316 | -105.6243
16.9353 | -105.6253
16.939 | -105.6265
16.9427 | -105.63
16.9464 | -105.6347
16.9501 | -105.6391
16.9538 | -105.6406
16.9576 | -105.6433
16.9613 | -105.6478
16.965 | -105.6502
16.9687 | -105.6551
16.9724 | -105.6633
16.9761 | -105.6717
16.9798 | -105.6819
16.9835 | -105.6915
16.9872 | -105.6983
16.991 | -105.7068
16.9947 | -105.7154
16.9984 | -105.7237
17.0021 | -105.732
17.0058 | -105.74
17.0095 | -105.7463
17.0132 | -105.75
17.0169 | -105.7559
17.0206 | -105.7606
17.0244 | -105.7616
17.0281 | -105.7593
17.0318 | -105.7577
17.0355 | -105.7566
17.0392 | -105.7563
17.0429 | -105.753
17.0466 | -105.7441
17.0503 | -105.7351
17.054 | -105.7238
17.0578 | -105.716
17.0615 | -105.7097
17.0652 | -105.7018
17.0689 | -105.6912
17.0726 | -105.6807
17.0763 | -105.6694
17.08 | -105.6567
17.0837 | -105.6453
17.0875 | -105.6316
17.0912 | -105.6197
17.0949 | -105.6124
17.0986 | -105.5994
17.1023 | -105.5856
17.106 | -105.5709
17.1097 | -105.5558
17.1134 | -105.5416
17.1171 | -105.5282
17.1209 | -105.5155
17.1246 | -105.5036
17.1283 | -105.4906
17.132 | -105.4854
17.1357 | -105.4755
17.1394 | -105.4665
17.1431 | -105.4553
17.1468 | -105.4444
17.1505 | -105.4342
17.1543 | -105.4246
17.158 | -105.4159
17.1617 | -105.4051
17.1654 | -105.3936
17.1691 | -105.3873
17.1728 | -105.3812
17.1765 | -105.3745
17.1802 | -105.368
17.1839 | -105.363
17.1877 | -105.3562
17.1914 | -105.3509
17.1951 | -105.3478
17.1988 | -105.3443
17.2025 | -105.3389
17.2062 | -105.3347
17.2099 | -105.3309
17.2136 | -105.3256
17.2174 | -105.3201
17.2211 | -105.3162
17.2248 | -105.313
17.2285 | -105.3118
17.2322 | -105.3111
17.2359 | -105.3085
17.2396 | -105.306
17.2433 | -105.3041
17.247 | -105.3027
17.2508 | -105.3024
17.2545 | -105.3021
17.2582 | -105.3039
17.2619 | -105.303
17.2656 | -105.3015
17.2693 | -105.3046
17.273 | -105.3101
17.2767 | -105.3134
17.2804 | -105.3165
17.2842 | -105.3193
17.2879 | -105.3244
17.2916 | -105.3295
17.2953 | -105.3328
17.299 | -105.3341
17.3027 | -105.3335
17.3064 | -105.3345
17.3101 | -105.3368
17.3138 | -105.3413
17.3176 | -105.3453
17.3213 | -105.3502
17.325 | -105.3556
17.3287 | -105.3585
17.3324 | -105.3625
17.3361 | -105.365
17.3398 | -105.3676
17.3435 | -105.3684
17.3473 | -105.3674
17.351 | -105.3683
17.3547 | -105.3693
17.3584 | -105.3672
17.3621 | -105.3646
17.3658 | -105.3621
17.3695 | -105.3583
17.3732 | -105.3544
17.3769 | -105.3482
17.3807 | -105.3397
17.3844 | -105.3323
17.3881 | -105.328
17.3918 | -105.324
17.3955 | -105.321
17.3992 | -105.3181
17.4029 | -105.3144
17.4066 | -105.3109
17.4103 | -105.3098
17.4141 | -105.3105
17.4178 | -105.3097
17.4215 | -105.3094
17.4252 | -105.3094
17.4289 | -105.3092
17.4326 | -105.31
17.4363 | -105.316
17.44 | -105.3216
17.4437 | -105.3242
17.4475 | -105.3288
17.4512 | -105.3368
17.4549 | -105.3427
17.4586 | -105.348
17.4623 | -105.3557
17.466 | -105.3649
17.4697 | -105.3703
17.4734 | -105.3785
17.4772 | -105.3859
17.4809 | -105.3941
17.4846 | -105.403
17.4883 | -105.4113
17.492 | -105.4181
17.4957 | -105.425
17.4994 | -105.4314
17.5031 | -105.4355
17.5068 | -105.4382
17.5106 | -105.4412
17.5143 | -105.4444
17.518 | -105.4448
17.5217 | -105.4424
17.5254 | -105.4393
17.5291 | -105.4376
17.5328 | -105.4369
17.5365 | -105.4356
17.5402 | -105.4328
17.544 | -105.4282
17.5477 | -105.4261
17.5514 | -105.4228
17.5551 | -105.4157
17.5588 | -105.4103
17.5625 | -105.4033
17.5662 | -105.3951
17.5699 | -105.3884
17.5736 | -105.3814
17.5774 | -105.3755
17.5811 | -105.3691
17.5848 | -105.3637
17.5885 | -105.3543
17.5922 | -105.3482
17.5959 | -105.3449
17.5996 | -105.3372
17.6033 | -105.3309
17.6071 | -105.3261
17.6108 | -105.3202
17.6145 | -105.3127
17.6182 | -105.3089
17.6219 | -105.3054
17.6256 | -105.3004
17.6293 | -105.2969
17.633 | -105.2926
17.6367 | -105.2865
17.6405 | -105.2812
17.6442 | -105.275
17.6479 | -105.2684
17.6516 | -105.2613
17.6553 | -105.2571
17.659 | -105.2519
17.6627 | -105.2462
17.6664 | -105.2428
17.6701 | -105.2372
17.6739 | -105.2339
17.6776 | -105.2314
17.6813 | -105.2314
17.685 | -105.2306
17.6887 | -105.2276
17.6924 | -105.2239
17.6961 | -105.2205
17.6998 | -105.2198
17.7035 | -105.2192
17.7073 | -105.2195
17.711 | -105.2193
17.7147 | -105.2189
17.7184 | -105.2176
17.7221 | -105.2193
17.7258 | -105.2203
17.7295 | -105.2223
17.7332 | -105.2247
17.737 | -105.2274
17.7407 | -105.2296
17.7444 | -105.2323
17.7481 | -105.2355
17.7518 | -105.24
17.7555 | -105.2461
17.7592 | -105.2501
17.7629 | -105.2555
17.7666 | -105.2607
17.7704 | -105.2645
17.7741 | -105.269
17.7778 | -105.2746
17.7815 | -105.2801
17.7852 | -105.2846
17.7889 | -105.2882
17.7926 | -105.2915
17.7963 | -105.2958
17.8 | -105.2994
17.8038 | -105.3037
17.8075 | -105.3081
17.8112 | -105.3101
17.8149 | -105.3139
17.8186 | -105.3175
17.8223 | -105.3193
17.826 | -105.3236
17.8297 | -105.3276
17.8334 | -105.3313
17.8372 | -105.3336
17.8409 | -105.3368
17.8446 | -105.3411
17.8483 | -105.3452
17.852 | -105.349
17.8557 | -105.3535
17.8594 | -105.3573
17.8631 | -105.3592
17.8669 | -105.3641
17.8706 | -105.367
17.8743 | -105.3703
17.878 | -105.378
17.8817 | -105.3854
17.8854 | -105.3937
17.8891 | -105.4024
17.8928 | -105.4118
17.8965 | -105.4216
17.9003 | -105.4301
17.904 | -105.4405
17.9077 | -105.448
17.9114 | -105.4564
17.9151 | -105.4683
17.9188 | -105.4797
17.9225 | -105.4896
17.9262 | -105.4992
17.9299 | -105.509
17.9337 | -105.5184
17.9374 | -105.5266
17.9411 | -105.5339
17.9448 | -105.5406
17.9485 | -105.5502
17.9522 | -105.5618
17.9559 | -105.5702
17.9596 | -105.575
17.9633 | -105.5811
17.9671 | -105.5896
17.9708 | -105.5944
17.9745 | -105.5967
17.9782 | -105.5967
17.9819 | -105.5961
17.9856 | -105.5959
17.9893 | -105.5952
17.993 | -105.5951
17.9968 | -105.593
18.0005 | -105.5896
18.0042 | -105.5844
18.0079 | -105.578
18.0116 | -105.5684
18.0153 | -105.5585
18.019 | -105.5498
18.0227 | -105.5426
18.0264 | -105.5333
18.0302 | -105.5213
18.0339 | -105.5077
18.0376 | -105.4943
18.0413 | -105.4817
18.045 | -105.4672
18.0487 | -105.4564
18.0524 | -105.4449
18.0561 | -105.432
18.0598 | -105.4182
18.0636 | -105.4058
18.0673 | -105.3957
18.071 | -105.3857
18.0747 | -105.3767
18.0784 | -105.3693
18.0821 | -105.359
18.0858 | -105.3503
18.0895 | -105.3424
18.0933 | -105.338
18.097 | -105.3326
18.1007 | -105.3275
18.1044 | -105.3239
18.1081 | -105.322
18.1118 | -105.3209
18.1155 | -105.3217
18.1192 | -105.3226
18.1229 | -105.323
18.1267 | -105.3264
18.1304 | -105.3293
18.1341 | -105.3315
18.1378 | -105.3361
18.1415 | -105.341
18.1452 | -105.3464
18.1489 | -105.3522
18.1526 | -105.3597
18.1563 | -105.3688
18.1601 | -105.3775
18.1638 | -105.3861
18.1675 | -105.3936
18.1712 | -105.3972
18.1749 | -105.3986
18.1786 | -105.4033
18.1823 | -105.4112
18.186 | -105.4157
18.1897 | -105.4185
18.1935 | -105.421
18.1972 | -105.4225
18.2009 | -105.4195
18.2046 | -105.4107
18.2083 | -105.4066
18.212 | -105.4015
18.2157 | -105.3961
18.2194 | -105.3922
18.2232 | -105.3867
18.2269 | -105.3792
18.2306 | -105.3721
18.2343 | -105.3653
18.238 | -105.3571
18.2417 | -105.3474
18.2454 | -105.3357
18.2491 | -105.3236
18.2528 | -105.3107
18.2566 | -105.3009
18.2603 | -105.2907
18.264 | -105.2793
18.2677 | -105.2666
18.2714 | -105.2554
18.2751 | -105.2439
18.2788 | -105.2301
18.2825 | -105.215
18.2862 | -105.2023
18.29 | -105.1926
18.2937 | -105.1826
18.2974 | -105.1701
18.3011 | -105.1575
18.3048 | -105.1493
18.3085 | -105.1407
18.3122 | -105.1306
18.3159 | -105.1192
18.3196 | -105.1093
18.3234 | -105.0992
18.3271 | -105.0886
18.3308 | -105.0808
18.3345 | -105.0747
18.3382 | -105.0683
18.3419 | -105.0649
18.3456 | -105.0605
18.3493 | -105.0558
18.353 | -105.0495
18.3568 | -105.046
18.3605 | -105.0428
18.3642 | -105.0395
18.3679 | -105.0383
18.3716 | -105.0371
18.3753 | -105.0349
18.379 | -105.034
18.3827 | -105.0334
18.3865 | -105.0334
18.3902 | -105.0359
18.3939 | -105.037
18.3976 | -105.038
18.4013 | -105.0405
18.405 | -105.045
18.4087 | -105.0513
18.4124 | -105.0564
18.4161 | -105.0612
18.4199 | -105.0653
18.4236 | -105.0704
18.4273 | -105.0786
18.431 | -105.0859
18.4347 | -105.0953
18.4384 | -105.1032
18.4421 | -105.1065
18.4458 | -105.1124
18.4495 | -105.118
18.4533 | -105.1217
18.457 | -105.1265
18.4607 | -105.1342
18.4644 | -105.1399
18.4681 | -105.1429
18.4718 | -105.1457
18.4755 | -105.1483
18.4792 | -105.1525
18.483 | -105.1568
18.4867 | -105.1596
18.4904 | -105.1608
18.4941 | -105.1609
18.4978 | -105.163
18.5015 | -105.1663
18.5052 | -105.1682
18.5089 | -105.1679
18.5126 | -105.1672
18.5164 | -105.1678
18.5201 | -105.1686
18.5238 | -105.1708
18.5275 | -105.1713
18.5312 | -105.1727
18.5349 | -105.1729
18.5386 | -105.1715
18.5423 | -105.1688
18.546 | -105.1676
18.5498 | -105.165
18.5535 | -105.1593
18.5572 | -105.1565
18.5609 | -105.155
18.5646 | -105.1528
18.5683 | -105.1489
18.572 | -105.1439
18.5757 | -105.1419
18.5794 | -105.14
18.5832 | -105.1362
18.5869 | -105.1319
18.5906 | -105.1293
18.5943 | -105.1244
18.598 | -105.1234
18.6017 | -105.1212
18.6054 | -105.118
18.6091 | -105.1166
18.6129 | -105.117
18.6166 | -105.1175
18.6203 | -105.1165
18.624 | -105.1154
18.6277 | -105.1147
18.6314 | -105.1151
18.6351 | -105.1159
18.6388 | -105.119
18.6425 | -105.1231
18.6463 | -105.1266
18.65 | -105.1277
18.6537 | -105.1274
18.6574 | -105.1313
18.6611 | -105.1369
18.6648 | -105.1418
18.6685 | -105.1436
18.6722 | -105.1459
18.6759 | -105.1511
18.6797 | -105.1589
18.6834 | -105.1661
18.6871 | -105.1746
18.6908 | -105.1827
18.6945 | -105.1917
18.6982 | -105.2008
18.7019 | -105.2109
18.7056 | -105.2242
18.7093 | -105.2364
18.7131 | -105.2501
18.7168 | -105.265
18.7205 | -105.2778
18.7242 | -105.2893
18.7279 | -105.3006
18.7316 | -105.3132
18.7353 | -105.3268
18.739 | -105.3395
18.7428 | -105.3513
18.7465 | -105.3618
18.7502 | -105.3711
18.7539 | -105.3822
18.7576 | -105.3936
18.7613 | -105.4058
18.765 | -105.4143
18.7687 | -105.4185
18.7724 | -105.4234
18.7762 | -105.4307
18.7799 | -105.4354
18.7836 | -105.4377
18.7873 | -105.4378
18.791 | -105.4358
18.7947 | -105.4335
18.7984 | -105.4294
18.8021 | -105.4249
18.8058 | -105.4202
18.8096 | -105.4138
18.8133 | -105.4056
18.817 | -105.3972
18.8207 | -105.3877
18.8244 | -105.3724
18.8281 | -105.3559
18.8318 | -105.3398
18.8355 | -105.3243
18.8392 | -105.309
18.843 | -105.292
18.8467 | -105.2775
18.8504 | -105.2645
18.8541 | -105.2497
18.8578 | -105.2307
18.8615 | -105.2131
18.8652 | -105.1954
18.8689 | -105.1791
18.8727 | -105.1634
18.8764 | -105.149
18.8801 | -105.1356
18.8838 | -105.1231
18.8875 | -105.1105
18.8912 | -105.0972
18.8949 | -105.0869
18.8986 | -105.0762
18.9023 | -105.0677
18.9061 | -105.0615
18.9098 | -105.0533
18.9135 | -105.0483
18.9172 | -105.0451
18.9209 | -105.0424
18.9246 | -105.0381
18.9283 | -105.0348
18.932 | -105.0351
18.9357 | -105.0375
18.9395 | -105.0386
18.9432 | -105.0378
18.9469 | -105.0392
18.9506 | -105.0441
18.9543 | -105.0485
18.958 | -105.0497
18.9617 | -105.0516
18.9654 | -105.0561
18.9691 | -105.0601
18.9729 | -105.0628
18.9766 | -105.0667
18.9803 | -105.0705
18.984 | -105.0749
18.9877 | -105.0784
18.9914 | -105.0828
18.9951 | -105.0852
18.9988 | -105.0853
19.0026 | -105.0863
19.0063 | -105.0868
19.01 | -105.0865
19.0137 | -105.0852
19.0174 | -105.0837
19.0211 | -105.083
19.0248 | -105.079
19.0285 | -105.076
19.0322 | -105.0729
19.036 | -105.0702
19.0397 | -105.0686
19.0434 | -105.0644
19.0471 | -105.0558
19.0508 | -105.0476
19.0545 | -105.0392
19.0582 | -105.0305
19.0619 | -105.0228
19.0656 | -105.0153
19.0694 | -105.0058
19.0731 | -104.9943
19.0768 | -104.9844
19.0805 | -104.9745
19.0842 | -104.9635
19.0879 | -104.9534
19.0916 | -104.9439
19.0953 | -104.936
19.099 | -104.9265
19.1028 | -104.9175
19.1065 | -104.9089
19.1102 | -104.9
19.1139 | -104.8907
19.1176 | -104.8793
19.1213 | -104.867
19.125 | -104.8565
19.1287 | -104.8495
19.1325 | -104.8409
19.1362 | -104.8301
19.1399 | -104.8212
19.1436 | -104.8138
19.1473 | -104.8076
19.151 | -104.7997
19.1547 | -104.7918
19.1584 | -104.7862
19.1621 | -104.7826
19.1659 | -104.7754
19.1696 | -104.7688
19.1733 | -104.7604
19.177 | -104.7534
19.1807 | -104.7495
19.1844 | -104.7422
19.1881 | -104.7353
19.1918 | -104.7319
19.1955 | -104.7273
19.1993 | -104.722
19.203 | -104.715
19.2067 | -104.7085
19.2104 | -104.7027
19.2141 | -104.6982
19.2178 | -104.6907
19.2215 | -104.6821
19.2252 | -104.673
19.2289 | -104.6661
19.2327 | -104.6612
19.2364 | -104.6529
19.2401 | -104.6427
19.2438 | -104.6363
19.2475 | -104.6298
19.2512 | -104.6236
19.2549 | -104.617
19.2586 | -104.6125
19.2624 | -104.6065
19.2661 | -104.6012
19.2698 | -104.5972
19.2735 | -104.5931
19.2772 | -104.5882
19.2809 | -104.5822
19.2846 | -104.5783
19.2883 | -104.5752
19.292 | -104.5699
19.2958 | -104.5637
19.2995 | -104.5603
19.3032 | -104.5573
19.3069 | -104.5529
19.3106 | -104.5494
19.3143 | -104.5477
19.318 | -104.5463
19.3217 | -104.5469
19.3254 | -104.5479
19.3292 | -104.5465
19.3329 | -104.5462
19.3366 | -104.5462
19.3403 | -104.5468
19.344 | -104.549
19.3477 | -104.5514
19.3514 | -104.551
19.3551 | -104.5521
19.3588 | -104.553
19.3626 | -104.552
19.3663 | -104.5538
19.37 | -104.5562
19.3737 | -104.5596
19.3774 | -104.5603
19.3811 | -104.5602
19.3848 | -104.5639
19.3885 | -104.5683
19.3923 | -104.5724
19.396 | -104.5771
19.3997 | -104.5804
19.4034 | -104.5851
19.4071 | -104.5914
19.4108 | -104.5958
19.4145 | -104.5992
19.4182 | -104.6019
19.4219 | -104.602
19.4257 | -104.6048
19.4294 | -104.6074
19.4331 | -104.6092
19.4368 | -104.6097
19.4405 | -104.6118
19.4442 | -104.6148
19.4479 | -104.6176
19.4516 | -104.618
19.4553 | -104.6192
19.4591 | -104.6209
19.4628 | -104.6232
19.4665 | -104.6235
19.4702 | -104.624
19.4739 | -104.6241
19.4776 | -104.6245
19.4813 | -104.6253
19.485 | -104.6257
19.4887 | -104.6258
19.4925 | -104.6282
19.4962 | -104.6299
19.4999 | -104.6313
19.5036 | -104.6324
19.5073 | -104.6316
19.511 | -104.6327
19.5147 | -104.6353
19.5184 | -104.6357
19.5222 | -104.6349
19.5259 | -104.6355
19.5296 | -104.6378
19.5333 | -104.6392
19.537 | -104.6397
19.5407 | -104.6392
19.5444 | -104.6395
19.5481 | -104.6397
19.5518 | -104.6371
19.5556 | -104.6345
19.5593 | -104.6342
19.563 | -104.6325
19.5667 | -104.6316
19.5704 | -104.6323
19.5741 | -104.6323
19.5778 | -104.6301
19.5815 | -104.6264
19.5852 | -104.6244
19.589 | -104.6221
19.5927 | -104.6185
19.5964 | -104.6156
19.6001 | -104.6134
19.6038 | -104.6114
19.6075 | -104.6105
19.6112 | -104.6069
19.6149 | -104.6049
19.6186 | -104.6035
19.6224 | -104.6023
19.6261 | -104.6034
19.6298 | -104.605
19.6335 | -104.6036
19.6372 | -104.6043
19.6409 | -104.6072
19.6446 | -104.6081
19.6483 | -104.6098
19.6521 | -104.6117
19.6558 | -104.616
19.6595 | -104.6194
19.6632 | -104.6192
19.6669 | -104.6218
19.6706 | -104.6231
19.6743 | -104.6237
19.678 | -104.626
19.6817 | -104.6239
19.6855 | -104.6219
19.6892 | -104.623
19.6929 | -104.6234
19.6966 | -104.621
19.7003 | -104.6178
19.704 | -104.6132
19.7077 | -104.6107
19.7114 | -104.6063
19.7151 | -104.6015
19.7189 | -104.5956
19.7226 | -104.5888
19.7263 | -104.5838
19.73 | -104.5786
19.7337 | -104.571
19.7374 | -104.5631
19.7411 | -104.5542
19.7448 | -104.5463
19.7486 | -104.5414
19.7523 | -104.5353
19.756 | -104.5279
19.7597 | -104.5215
19.7634 | -104.5148
19.7671 | -104.5087
19.7708 | -104.5022
19.7745 | -104.4984
19.7782 | -104.4935
19.782 | -104.489
19.7857 | -104.4844
19.7894 | -104.4779
19.7931 | -104.4742
19.7968 | -104.4695
19.8005 | -104.4669
19.8042 | -104.4636
19.8079 | -104.466
19.8116 | -104.4642
19.8154 | -104.4638
19.8191 | -104.4652
19.8228 | -104.4666
19.8265 | -104.4681
19.8302 | -104.4703
19.8339 | -104.4746
19.8376 | -104.4795
19.8413 | -104.4827
19.845 | -104.4837
19.8488 | -104.4852
19.8525 | -104.4886
19.8562 | -104.4942
19.8599 | -104.5017
19.8636 | -104.5102
19.8673 | -104.5175
19.871 | -104.5238
19.8747 | -104.5306
19.8785 | -104.5366
19.8822 | -104.544
19.8859 | -104.553
19.8896 | -104.562
19.8933 | -104.5738
19.897 | -104.583
19.9007 | -104.5917
19.9044 | -104.5991
19.9081 | -104.6042
19.9119 | -104.612
19.9156 | -104.6225
19.9193 | -104.6303
19.923 | -104.6363
19.9267 | -104.6417
19.9304 | -104.6457
19.9341 | -104.6495
19.9378 | -104.6542
19.9415 | -104.658
19.9453 | -104.6627
19.949 | -104.663
19.9527 | -104.6624
19.9564 | -104.6619
19.9601 | -104.6606
19.9638 | -104.6597
19.9675 | -104.6579
19.9712 | -104.6542
19.9749 | -104.6501
19.9787 | -104.6474
19.9824 | -104.6434
19.9861 | -104.6391
19.9898 | -104.6332
19.9935 | -104.6274
19.9972 | -104.6216
20.0009 | -104.6171
20.0046 | -104.6114
20.0084 | -104.6046
20.0121 | -104.6001
20.0158 | -104.5956
20.0195 | -104.5912
20.0232 | -104.5865
20.0269 | -104.5821
20.0306 | -104.577
20.0343 | -104.5723
20.038 | -104.5678
20.0418 | -104.563
20.0455 | -104.561
20.0492 | -104.5591
20.0529 | -104.5551
20.0566 | -104.5518
20.0603 | -104.5524
20.064 | -104.5524
20.0677 | -104.5535
20.0714 | -104.5567
20.0752 | -104.5552
20.0789 | -104.554
20.0826 | -104.554
20.0863 | -104.5547
20.09 | -104.5558
20.0937 | -104.5555
20.0974 | -104.5541
20.1011 | -104.5521
20.1048 | -104.5515
20.1086 | -104.5513
20.1123 | -104.5496
20.116 | -104.5477
20.1197 | -104.5457
20.1234 | -104.5416
20.1271 | -104.5358
20.1308 | -104.5301
20.1345 | -104.5253
20.1383 | -104.5189
20.142 | -104.5105
20.1457 | -104.5027
20.1494 | -104.4951
20.1531 | -104.4863
20.1568 | -104.4772
20.1605 | -104.4648
20.1642 | -104.4545
20.1679 | -104.4439
20.1717 | -104.4305
20.1754 | -104.4168
20.1791 | -104.4026
20.1828 | -104.3897
20.1865 | -104.3766
20.1902 | -104.3662
20.1939 | -104.3557
20.1976 | -104.3424
20.2013 | -104.3263
20.2051 | -104.3111
20.2088 | -104.2979
20.2125 | -104.2855
20.2162 | -104.2739
20.2199 | -104.2614
20.2236 | -104.2472
20.2273 | -104.234
20.231 | -104.221
20.2347 | -104.2113
20.2385 | -104.204
20.2422 | -104.1963
20.2459 | -104.1886
20.2496 | -104.1824
20.2533 | -104.1764
20.257 | -104.1716
20.2607 | -104.1683
20.2644 | -104.1671
20.2682 | -104.1677
20.2719 | -104.1676
20.2756 | -104.1684
20.2793 | -104.1722
20.283 | -104.1771
20.2867 | -104.181
20.2904 | -104.1876
20.2941 | -104.1951
20.2978 | -104.2026
20.3016 | -104.2143
20.3053 | -104.2296
20.309 | -104.2453
20.3127 | -104.2593
20.3164 | -104.274
20.3201 | -104.2913
20.3238 | -104.3093
20.3275 | -104.3286
20.3312 | -104.347
20.335 | -104.3631
20.3387 | -104.3815
20.3424 | -104.3995
20.3461 | -104.4133
20.3498 | -104.4243
20.3535 | -104.4388
20.3572 | -104.4537
20.3609 | -104.4661
20.3646 | -104.4805
20.3684 | -104.4943
20.3721 | -104.5036
20.3758 | -104.5132
20.3795 | -104.5212
20.3832 | -104.5278
20.3869 | -104.5335
20.3906 | -104.5383
20.3943 | -104.5396
20.3981 | -104.5392
20.4018 | -104.5392
20.4055 | -104.5388
20.4092 | -104.537
20.4129 | -104.5358
20.4166 | -104.5334
20.4203 | -104.5297
20.424 | -104.5238
20.4277 | -104.5161
20.4315 | -104.5084
20.4352 | -104.4985
20.4389 | -104.4891
20.4426 | -104.4796
20.4463 | -104.4704
20.45 | -104.4629
20.4537 | -104.4544
20.4574 | -104.4436
20.4611 | -104.4353
20.4649 | -104.4272
20.4686 | -104.4186
20.4723 | -104.4096
20.476 | -104.4015
20.4797 | -104.3918
20.4834 | -104.385
20.4871 | -104.3782
20.4908 | -104.3724
20.4945 | -104.3673
20.4983 | -104.3616
20.502 | -104.3528
20.5057 | -104.3438
20.5094 | -104.3393
20.5131 | -104.3357
20.5168 | -104.3331
20.5205 | -104.3298
20.5242 | -104.3247
20.528 | -104.3202
20.5317 | -104.3163
20.5354 | -104.312
20.5391 | -104.3066
20.5428 | -104.3011
20.5465 | -104.2963
20.5502 | -104.2899
20.5539 | -104.2827
20.5576 | -104.2777
20.5614 | -104.2744
20.5651 | -104.2682
20.5688 | -104.2602
20.5725 | -104.2536
20.5762 | -104.2489
20.5799 | -104.242
20.5836 | -104.2344
20.5873 | -104.2282
20.591 | -104.2204
20.5948 | -104.2095
20.5985 | -104.1997
20.6022 | -104.1915
20.6059 | -104.1857
20.6096 | -104.1769
20.6133 | -104.1684
20.617 | -104.1569
20.6207 | -104.1467
20.6244 | -104.1381
20.6282 | -104.128
20.6319 | -104.1184
20.6356 | -104.1085
20.6393 | -104.0989
20.643 | -104.0903
20.6467 | -104.0805
20.6504 | -104.0691
20.6541 | -104.0591
20.6579 | -104.0474
20.6616 | -104.0369
20.6653 | -104.0284
20.669 | -104.0182
20.6727 | -104.0084
20.6764 | -103.9989
20.6801 | -103.9903
20.6838 | -103.9812
20.6875 | -103.9707
20.6913 | -103.9606
20.695 | -103.9498
20.6987 | -103.9396
20.7024 | -103.9301
20.7061 | -103.9209
20.7098 | -103.9136
20.7135 | -103.9074
20.7172 | -103.8997
20.7209 | -103.8913
20.7247 | -103.8826
20.7284 | -103.8764
20.7321 | -103.8693
20.7358 | -103.8632
20.7395 | -103.8579
20.7432 | -103.8535
20.7469 | -103.8479
20.7506 | -103.8426
20.7543 | -103.8368
20.7581 | -103.8301
20.7618 | -103.824
20.7655 | -103.8184
20.7692 | -103.8133
20.7729 | -103.8088
20.7766 | -103.8016
20.7803 | -103.7944
20.784 | -103.7884
20.7878 | -103.7847
20.7915 | -103.78
20.7952 | -103.777
20.7989 | -103.7726
20.8026 | -103.7657
20.8063 | -103.7577
20.81 | -103.7487
20.8137 | -103.7404
20.8174 | -103.7323
20.8212 | -103.723
20.8249 | -103.7141
20.8286 | -103.7052
20.8323 | -103.6988
20.836 | -103.6927
20.8397 | -103.6858
20.8434 | -103.6795
20.8471 | -103.6731
20.8508 | -103.6642
20.8546 | -103.6553
20.8583 | -103.649
20.862 | -103.644
20.8657 | -103.6366
20.8694 | -103.6272
20.8731 | -103.619
20.8768 | -103.613
20.8805 | -103.6073
20.8843 | -103.6005
20.888 | -103.593
20.8917 | -103.5872
20.8954 | -103.5824
20.8991 | -103.5786
20.9028 | -103.5728
20.9065 | -103.569
20.9102 | -103.5667
20.9139 | -103.5653
20.9177 | -103.5629
20.9214 | -103.5627
20.9251 | -103.5635
20.9288 | -103.5628
20.9325 | -103.5644
20.9362 | -103.5685
20.9399 | -103.5749
20.9436 | -103.58
20.9473 | -103.586
20.9511 | -103.5923
20.9548 | -103.6021
20.9585 | -103.6142
20.9622 | -103.6253
20.9659 | -103.6359
20.9696 | -103.6502
20.9733 | -103.6662
20.977 | -103.6829
20.9807 | -103.6995
20.9845 | -103.7185
20.9882 | -103.7378
20.9919 | -103.757
20.9956 | -103.7744
20.9993 | -103.7911
21.003 | -103.8078
21.0067 | -103.8246
21.0104 | -103.842
21.0142 | -103.8616
21.0179 | -103.8817
21.0216 | -103.901
21.0253 | -103.9183
21.029 | -103.9347
21.0327 | -103.9525
21.0364 | -103.9701
21.0401 | -103.984
21.0438 | -103.9986
21.0476 | -104.0147
21.0513 | -104.0289
21.055 | -104.0438
21.0587 | -104.0568
21.0624 | -104.069
21.0661 | -104.0801
21.0698 | -104.0907
21.0735 | -104.1027
21.0772 | -104.1137
21.081 | -104.1236
21.0847 | -104.1306
21.0884 | -104.1384
21.0921 | -104.1452
21.0958 | -104.1516
21.0995 | -104.1573
21.1032 | -104.1629
21.1069 | -104.1678
21.1106 | -104.1707
21.1144 | -104.1744
21.1181 | -104.1757
21.1218 | -104.1746
21.1255 | -104.1748
21.1292 | -104.1751
21.1329 | -104.1757
21.1366 | -104.1755
21.1403 | -104.1743
21.1441 | -104.1723
21.1478 | -104.1727
21.1515 | -104.1739
21.1552 | -104.1738
21.1589 | -104.1714
21.1626 | -104.1709
21.1663 | -104.1742
21.17 | -104.1762
21.1737 | -104.1765
21.1775 | -104.1775
21.1812 | -104.1797
21.1849 | -104.1826
21.1886 | -104.1858
21.1923 | -104.1874
21.196 | -104.1875
21.1997 | -104.1864
21.2034 | -104.1879
21.2071 | -104.1905
21.2109 | -104.1924
21.2146 | -104.1935
21.2183 | -104.194
21.222 | -104.1941
21.2257 | -104.1946
21.2294 | -104.1944
21.2331 | -104.1924
21.2368 | -104.1907
21.2405 | -104.1878
21.2443 | -104.1861
21.248 | -104.1842
21.2517 | -104.1809
21.2554 | -104.1763
21.2591 | -104.1681
21.2628 | -104.162
21.2665 | -104.1576
21.2702 | -104.1489
21.2739 | -104.1396
21.2777 | -104.1305
21.2814 | -104.1225
21.2851 | -104.1135
21.2888 | -104.1043
21.2925 | -104.0958
21.2962 | -104.0879
21.2999 | -104.08
21.3036 | -104.0715
21.3074 | -104.0613
21.3111 | -104.0499
21.3148 | -104.0392
21.3185 | -104.0288
21.3222 | -104.0185
21.3259 | -104.0083
21.3296 | -103.9987
21.3333 | -103.988
21.337 | -103.9758
21.3408 | -103.9667
21.3445 | -103.9569
21.3482 | -103.9469
21.3519 | -103.938
21.3556 | -103.9304
21.3593 | -103.9236
21.363 | -103.9159
21.3667 | -103.9069
21.3704 | -103.8997
21.3742 | -103.892
21.3779 | -103.8848
21.3816 | -103.8764
21.3853 | -103.8697
21.389 | -103.8628
21.3927 | -103.8574
21.3964 | -103.853
21.4001 | -103.8466
21.4039 | -103.8418
21.4076 | -103.8375
21.4113 | -103.8335
21.415 | -103.8291
21.4187 | -103.823
21.4224 | -103.8174
21.4261 | -103.8128
21.4298 | -103.8061
21.4335 | -103.7988
21.4373 | -103.7919
21.441 | -103.7865
21.4447 | -103.7814
21.4484 | -103.7774
21.4521 | -103.7734
21.4558 | -103.7685
21.4595 | -103.7651
21.4632 | -103.7626
21.4669 | -103.7603
21.4707 | -103.7551
21.4744 | -103.7532
21.4781 | -103.7536
21.4818 | -103.7539
21.4855 | -103.7524
21.4892 | -103.7506
21.4929 | -103.749
21.4966 | -103.7507
21.5003 | -103.7545
21.5041 | -103.76
21.5078 | -103.7632
21.5115 | -103.7667
21.5152 | -103.7696
21.5189 | -103.7754
21.5226 | -103.7817
21.5263 | -103.7864
21.53 | -103.791
21.5338 | -103.7979
21.5375 | -103.8047
21.5412 | -103.8089
21.5449 | -103.8121
21.5486 | -103.8173
21.5523 | -103.8221
21.556 | -103.8288
21.5597 | -103.8354
21.5634 | -103.8437
21.5672 | -103.8507
21.5709 | -103.8553
21.5746 | -103.8601
21.5783 | -103.8651
21.582 | -103.8688
21.5857 | -103.8699
21.5894 | -103.8709
21.5931 | -103.8733
21.5968 | -103.8741
21.6006 | -103.8745
21.6043 | -103.8744
21.608 | -103.8741
21.6117 | -103.8749
21.6154 | -103.8745
21.6191 | -103.8715
21.6228 | -103.8683
21.6265 | -103.8648
21.6302 | -103.861
21.634 | -103.8571
21.6377 | -103.8519
21.6414 | -103.8486
21.6451 | -103.8427
21.6488 | -103.836
21.6525 | -103.8293
21.6562 | -103.8235
21.6599 | -103.8171
21.6637 | -103.8088
21.6674 | -103.8021
21.6711 | -103.7955
21.6748 | -103.7864
21.6785 | -103.7777
21.6822 | -103.7701
21.6859 | -103.7636
21.6896 | -103.7558
21.6933 | -103.7488
21.6971 | -103.7422
21.7008 | -103.7355
21.7045 | -103.7299
21.7082 | -103.7235
21.7119 | -103.7162
21.7156 | -103.709
21.7193 | -103.7026
21.723 | -103.6987
21.7267 | -103.6953
21.7305 | -103.6931
21.7342 | -103.6906
21.7379 | -103.6896
21.7416 | -103.6887
21.7453 | -103.6858
21.749 | -103.6839
21.7527 | -103.684
21.7564 | -103.6831
21.7601 | -103.6823
21.7639 | -103.6809
21.7676 | -103.6799
21.7713 | -103.6808
21.775 | -103.6813
21.7787 | -103.6824
21.7824 | -103.6833
21.7861 | -103.684
21.7898 | -103.6865
21.7936 | -103.688
21.7973 | -103.691
21.801 | -103.695
21.8047 | -103.6981
21.8084 | -103.7003
21.8121 | -103.7033
21.8158 | -103.7082
21.8195 | -103.7118
21.8232 | -103.7155
21.827 | -103.7191
21.8307 | -103.7222
21.8344 | -103.7271
21.8381 | -103.7322
21.8418 | -103.7366
21.8455 | -103.7412
21.8492 | -103.7481
21.8529 | -103.7544
21.8566 | -103.7597
21.8604 | -103.7642
21.8641 | -103.766
21.8678 | -103.7686
21.8715 | -103.7723
21.8752 | -103.7759
21.8789 | -103.7809
21.8826 | -103.7864
21.8863 | -103.792
21.89 | -103.7968
21.8938 | -103.8018
21.8975 | -103.8068
21.9012 | -103.8092
21.9049 | -103.8116
21.9086 | -103.8141
21.9123 | -103.8168
21.916 | -103.8172
21.9197 | -103.8175
21.9235 | -103.8179
21.9272 | -103.8185
21.9309 | -103.8186
21.9346 | -103.8178
21.9383 | -103.816
21.942 | -103.8131
21.9457 | -103.8128
21.9494 | -103.8107
21.9531 | -103.8086
21.9569 | -103.8072
21.9606 | -103.8062
21.9643 | -103.8042
21.968 | -103.8006
21.9717 | -103.7961
21.9754 | -103.794
21.9791 | -103.7916
21.9828 | -103.7884
21.9865 | -103.7857
21.9903 | -103.7828
21.994 | -103.779
21.9977 | -103.776
22.0014 | -103.7721
22.0051 | -103.7663
22.0088 | -103.7604
22.0125 | -103.7548
22.0162 | -103.7509
22.0199 | -103.7461
22.0237 | -103.741
22.0274 | -103.7373
22.0311 | -103.7336
22.0348 | -103.7275
22.0385 | -103.7217
22.0422 | -103.7162
22.0459 | -103.7097
22.0496 | -103.7014
22.0534 | -103.6927
22.0571 | -103.6872
22.0608 | -103.6798
22.0645 | -103.6715
22.0682 | -103.6656
22.0719 | -103.6613
22.0756 | -103.6564
22.0793 | -103.6517
22.083 | -103.6463
22.0868 | -103.6414
22.0905 | -103.6374
22.0942 | -103.6334
22.0979 | -103.6297
22.1016 | -103.6265
22.1053 | -103.6256
22.109 | -103.6241
22.1127 | -103.6226
22.1164 | -103.62
22.1202 | -103.6178
22.1239 | -103.6143
22.1276 | -103.6129
22.1313 | -103.6143
22.135 | -103.6182
22.1387 | -103.6212
22.1424 | -103.6248
22.1461 | -103.6274
22.1498 | -103.6308
22.1536 | -103.6341
22.1573 | -103.6386
22.161 | -103.6442
22.1647 | -103.6513
22.1684 | -103.6577
22.1721 | -103.6654
22.1758 | -103.6748
22.1795 | -103.6841
22.1833 | -103.6925
22.187 | -103.7027
22.1907 | -103.7144
22.1944 | -103.7242
22.1981 | -103.7315
22.2018 | -103.7419
22.2055 | -103.7553
22.2092 | -103.7676
22.2129 | -103.7782
22.2167 | -103.7891
22.2204 | -103.7982
22.2241 | -103.8071
22.2278 | -103.8179
22.2315 | -103.8252
22.2352 | -103.8324
22.2389 | -103.838
22.2426 | -103.8418
22.2463 | -103.8456
22.2501 | -103.8481
22.2538 | -103.85
22.2575 | -103.8491
22.2612 | -103.8471
22.2649 | -103.8447
22.2686 | -103.8418
22.2723 | -103.837
22.276 | -103.8302
22.2797 | -103.821
22.2835 | -103.8128
22.2872 | -103.8026
22.2909 | -103.7892
22.2946 | -103.7739
22.2983 | -103.7602
22.302 | -103.7458
22.3057 | -103.73
22.3094 | -103.7138
22.3132 | -103.6958
22.3169 | -103.6781
22.3206 | -103.6626
22.3243 | -103.647
22.328 | -103.6294
22.3317 | -103.6102
22.3354 | -103.5918
22.3391 | -103.5745
22.3428 | -103.5564
22.3466 | -103.5385
22.3503 | -103.5216
22.354 | -103.505
22.3577 | -103.4872
22.3614 | -103.4713
22.3651 | -103.46
22.3688 | -103.4487
22.3725 | -103.4363
22.3762 | -103.425
22.38 | -103.4164
22.3837 | -103.4076
22.3874 | -103.3993
22.3911 | -103.3933
22.3948 | -103.3879
22.3985 | -103.3848
22.4022 | -103.3815
22.4059 | -103.38
22.4096 | -103.3775
22.4134 | -103.3724
22.4171 | -103.3672
22.4208 | -103.3661
22.4245 | -103.3667
22.4282 | -103.3676
22.4319 | -103.3666
22.4356 | -103.3611
22.4393 | -103.3566
22.4431 | -103.3526
22.4468 | -103.3488
22.4505 | -103.3459
22.4542 | -103.3431
22.4579 | -103.3389
22.4616 | -103.3345
22.4653 | -103.3305
22.469 | -103.3224
22.4727 | -103.3105
22.4765 | -103.298
22.4802 | -103.2806
22.4839 | -103.2587
22.4876 | -103.2205
22.4913 | -103.2011
22.495 | -103.179
22.4987 | -103.1555
22.5024 | -103.1267
22.5061 | -103.0922
22.5099 | -103.0551
22.5136 | -103.0172
22.5173 | -102.9766
22.521 | -102.9342
22.5247 | -102.8928
22.5284 | -102.8507
22.5321 | -102.8096
22.5358 | -102.7687
22.5395 | -102.7274
22.5433 | -102.683
22.547 | -102.6387
22.5507 | -102.5958
22.5544 | -102.5526
22.5581 | -102.5102
22.5618 | -102.4669
22.5655 | -102.4245
22.5692 | -102.3847
22.573 | -102.3451
22.5767 | -102.3054
22.5804 | -102.2674
22.5841 | -102.2303
22.5878 | -102.196
22.5915 | -102.1641
22.5952 | -102.1355
22.5989 | -102.1104
22.6026 | -102.0863
22.6064 | -102.0671
22.6101 | -102.0519
22.6138 | -102.039
22.6175 | -102.027
22.6212 | -102.0179
22.6249 | -102.0118
22.6286 | -102.0087
22.6323 | -102.0087
22.636 | -102.0112
22.6398 | -102.0159
22.6435 | -102.0241
22.6472 | -102.0366
22.6509 | -102.0515
22.6546 | -102.0681
22.6583 | -102.0887
22.662 | -102.112
22.6657 | -102.1375
22.6695 | -102.1668
22.6732 | -102.1974
22.6769 | -102.2297
22.6806 | -102.2632
22.6843 | -102.2959
22.688 | -102.3307
22.6917 | -102.3681
22.6954 | -102.4075
22.6991 | -102.4462
22.7029 | -102.4846
22.7066 | -102.5241
22.7103 | -102.5643
22.714 | -102.6052
22.7177 | -102.6452
22.7214 | -102.6857
22.7251 | -102.7238
22.7288 | -102.7629
22.7325 | -102.8021
22.7363 | -102.8412
22.74 | -102.8743
22.7437 | -102.9047
22.7474 | -102.9365
22.7511 | -102.9694
22.7548 | -103.0007
22.7585 | -103.0288
22.7622 | -103.052
22.7659 | -103.0723
22.7697 | -103.0875
22.7734 | -103.1016
22.7771 | -103.1173
22.7808 | -103.131
22.7845 | -103.14
22.7882 | -103.1495
22.7919 | -103.1582
22.7956 | -103.1639
22.7994 | -103.1713
22.8031 | -103.1787
22.8068 | -103.1803
22.8105 | -103.18
22.8142 | -103.1772
22.8179 | -103.1761
22.8216 | -103.1714
22.8253 | -103.168
22.829 | -103.1684
22.8328 | -103.1683
22.8365 | -103.1671
22.8402 | -103.1678
22.8439 | -103.1651
22.8476 | -103.1568
22.8513 | -103.1491
22.855 | -103.1432
22.8587 | -103.1342
22.8624 | -103.1279
22.8662 | -103.1257
22.8699 | -103.1228
22.8736 | -103.1201
22.8773 | -103.1169
22.881 | -103.1149
22.8847 | -103.1175
22.8884 | -103.1179
22.8921 | -103.1166
22.8958 | -103.1189
22.8996 | -103.1217
22.9033 | -103.1252
22.907 | -103.129
22.9107 | -103.1327
22.9144 | -103.1378
22.9181 | -103.1421
22.9218 | -103.1448
22.9255 | -103.1493
22.9293 | -103.1541
22.933 | -103.1584
22.9367 | -103.1637
22.9404 | -103.1694
22.9441 | -103.1754
22.9478 | -103.1814
22.9515 | -103.1861
22.9552 | -103.1909
22.9589 | -103.1953
22.9627 | -103.2
22.9664 | -103.2043
22.9701 | -103.2088
22.9738 | -103.2114
22.9775 | -103.2135
22.9812 | -103.2139
22.9849 | -103.2137
22.9886 | -103.2127
22.9923 | -103.2082
22.9961 | -103.2063
22.9998 | -103.2064
23.0035 | -103.2065
23.0072 | -103.206
23.0109 | -103.2025
23.0146 | -103.1987
23.0183 | -103.1957
23.022 | -103.19
23.0257 | -103.1865
23.0295 | -103.1824
23.0332 | -103.1767
23.0369 | -103.1713
23.0406 | -103.1642
23.0443 | -103.1579
23.048 | -103.1528
23.0517 | -103.1494
23.0554 | -103.146
23.0592 | -103.1444
23.0629 | -103.142
23.0666 | -103.1403
23.0703 | -103.1386
23.074 | -103.134
23.0777 | -103.1297
23.0814 | -103.1273
23.0851 | -103.1248
23.0888 | -103.1243
23.0926 | -103.123
23.0963 | -103.1232
23.1 | -103.1209
23.1037 | -103.1167
23.1074 | -103.1155
23.1111 | -103.1167
23.1148 | -103.1193
23.1185 | -103.1224
23.1222 | -103.126
23.126 | -103.1302
23.1297 | -103.1356
23.1334 | -103.1403
23.1371 | -103.1431
23.1408 | -103.1465
23.1445 | -103.1478
23.1482 | -103.1528
23.1519 | -103.1612
23.1556 | -103.1675
23.1594 | -103.1747
23.1631 | -103.1821
23.1668 | -103.1876
23.1705 | -103.1941
23.1742 | -103.2033
23.1779 | -103.2096
23.1816 | -103.2125
23.1853 | -103.2202
23.1891 | -103.2272
23.1928 | -103.2333
23.1965 | -103.2398
23.2002 | -103.2452
23.2039 | -103.251
23.2076 | -103.2558
23.2113 | -103.2582
23.215 | -103.2619
23.2187 | -103.2663
23.2225 | -103.2681
23.2262 | -103.2696
23.2299 | -103.2722
23.2336 | -103.275
23.2373 | -103.2778
23.241 | -103.2791
23.2447 | -103.2796
23.2484 | -103.2811
23.2521 | -103.2804
23.2559 | -103.2791
23.2596 | -103.2788
23.2633 | -103.2778
23.267 | -103.2774
23.2707 | -103.2766
23.2744 | -103.2756
23.2781 | -103.2756
23.2818 | -103.2736
23.2855 | -103.2716
23.2893 | -103.2701
23.293 | -103.2677
23.2967 | -103.2648
23.3004 | -103.264
23.3041 | -103.2637
23.3078 | -103.2626
23.3115 | -103.2622
23.3152 | -103.2614
23.319 | -103.2616
23.3227 | -103.262
23.3264 | -103.2605
23.3301 | -103.2558
23.3338 | -103.2498
23.3375 | -103.2473
23.3412 | -103.244
23.3449 | -103.2385
23.3486 | -103.235
23.3524 | -103.2325
23.3561 | -103.2306
23.3598 | -103.2298
23.3635 | -103.2294
23.3672 | -103.2284
23.3709 | -103.2285
23.3746 | -103.2267
23.3783 | -103.2243
23.382 | -103.2225
23.3858 | -103.2207
23.3895 | -103.2178
23.3932 | -103.2139
23.3969 | -103.2111
23.4006 | -103.2096
23.4043 | -103.2091
23.408 | -103.2082
23.4117 | -103.2092
23.4154 | -103.2099
23.4192 | -103.2087
23.4229 | -103.2088
23.4266 | -103.2108
23.4303 | -103.215
23.434 | -103.2172
23.4377 | -103.2159
23.4414 | -103.2138
23.4451 | -103.213
23.4489 | -103.214
23.4526 | -103.2153
23.4563 | -103.2178
23.46 | -103.219
23.4637 | -103.222
23.4674 | -103.2266
23.4711 | -103.2285
23.4748 | -103.2337
23.4785 | -103.2378
23.4823 | -103.2384
23.486 | -103.2401
23.4897 | -103.2413
23.4934 | -103.2433
23.4971 | -103.2477
23.5008 | -103.2512
23.5045 | -103.2538
23.5082 | -103.2568
23.5119 | -103.26
23.5157 | -103.2601
23.5194 | -103.2607
23.5231 | -103.2611
23.5268 | -103.2623
23.5305 | -103.2647
23.5342 | -103.267
23.5379 | -103.2693
23.5416 | -103.2713
23.5453 | -103.2724
23.5491 | -103.2715
23.5528 | -103.2717
23.5565 | -103.2728
23.5602 | -103.2736
23.5639 | -103.2732
23.5676 | -103.2722
23.5713 | -103.271
23.575 | -103.2702
23.5788 | -103.271
23.5825 | -103.2696
23.5862 | -103.2693
23.5899 | -103.2691
23.5936 | -103.2686
23.5973 | -103.2699
23.601 | -103.2729
23.6047 | -103.2758
23.6084 | -103.2789
23.6122 | -103.2797
23.6159 | -103.2827
23.6196 | -103.2875
23.6233 | -103.2913
23.627 | -103.2947
23.6307 | -103.2996
23.6344 | -103.3051
23.6381 | -103.3101
23.6418 | -103.315
23.6456 | -103.3212
23.6493 | -103.3279
23.653 | -103.3355
23.6567 | -103.3418
23.6604 | -103.3489
23.6641 | -103.3598
23.6678 | -103.3727
23.6715 | -103.384
23.6752 | -103.3936
23.679 | -103.403
23.6827 | -103.4106
23.6864 | -103.4172
23.6901 | -103.4249
23.6938 | -103.4351
23.6975 | -103.4438
23.7012 | -103.4522
23.7049 | -103.4608
23.7087 | -103.4699
23.7124 | -103.4789
23.7161 | -103.4871
23.7198 | -103.4922
23.7235 | -103.4987
23.7272 | -103.507
23.7309 | -103.5153
23.7346 | -103.5206
23.7383 | -103.5261
23.7421 | -103.5308
23.7458 | -103.5349
23.7495 | -103.5386
23.7532 | -103.5427
23.7569 | -103.5459
23.7606 | -103.5478
23.7643 | -103.5497
23.768 | -103.5536
23.7717 | -103.5567
23.7755 | -103.5597
23.7792 | -103.5605
23.7829 | -103.5592
23.7866 | -103.5567
23.7903 | -103.5518
23.794 | -103.5495
23.7977 | -103.5466
23.8014 | -103.5434
23.8052 | -103.5415
23.8089 | -103.5375
23.8126 | -103.5327
23.8163 | -103.5281
23.82 | -103.5228
23.8237 | -103.517
23.8274 | -103.5106
23.8311 | -103.502
23.8348 | -103.4949
23.8386 | -103.4869
23.8423 | -103.4798
23.846 | -103.4725
23.8497 | -103.4661
23.8534 | -103.4582
23.8571 | -103.4484
23.8608 | -103.4397
23.8645 | -103.4307
23.8682 | -103.4215
23.872 | -103.4128
23.8757 | -103.4036
23.8794 | -103.3932
23.8831 | -103.3834
23.8868 | -103.3745
23.8905 | -103.3652
23.8942 | -103.3568
23.8979 | -103.3508
23.9016 | -103.3444
23.9054 | -103.3372
23.9091 | -103.33
23.9128 | -103.3244
23.9165 | -103.3198
23.9202 | -103.3162
23.9239 | -103.3116
23.9276 | -103.306
23.9313 | -103.3
23.9351 | -103.2936
23.9388 | -103.2864
23.9425 | -103.2802
23.9462 | -103.2746
23.9499 | -103.2718
23.9536 | -103.2682
23.9573 | -103.263
23.961 | -103.2575
23.9647 | -103.2541
23.9685 | -103.2495
23.9722 | -103.2461
23.9759 | -103.2442
23.9796 | -103.2409
23.9833 | -103.2373
23.987 | -103.2349
23.9907 | -103.2326
23.9944 | -103.2308
23.9981 | -103.2274
24.0019 | -103.2241
24.0056 | -103.2214
24.0093 | -103.2178
24.013 | -103.2155
24.0167 | -103.2155
24.0204 | -103.2146
24.0241 | -103.2137
24.0278 | -103.2108
24.0315 | -103.2059
24.0353 | -103.2042
24.039 | -103.2033
24.0427 | -103.2021
24.0464 | -103.2012
24.0501 | -103.2009
24.0538 | -103.2013
24.0575 | -103.2024
24.0612 | -103.2034
24.0649 | -103.204
24.0687 | -103.2048
24.0724 | -103.2068
24.0761 | -103.2088
24.0798 | -103.2122
24.0835 | -103.2155
24.0872 | -103.2174
24.0909 | -103.2199
24.0946 | -103.2279
24.0984 | -103.2316
24.1021 | -103.2351
24.1058 | -103.2389
24.1095 | -103.2416
24.1132 | -103.2464
24.1169 | -103.2532
24.1206 | -103.2602
24.1243 | -103.2671
24.128 | -103.2729
24.1318 | -103.2805
24.1355 | -103.2867
24.1392 | -103.2938
24.1429 | -103.3004
24.1466 | -103.3069
24.1503 | -103.314
24.154 | -103.32
24.1577 | -103.3266
24.1614 | -103.3331
24.1652 | -103.3412
24.1689 | -103.348
24.1726 | -103.3552
24.1763 | -103.362
24.18 | -103.3671
24.1837 | -103.3734
24.1874 | -103.3795
24.1911 | -103.3853
24.1948 | -103.3916
24.1986 | -103.3992
24.2023 | -103.4055
24.206 | -103.4113
24.2097 | -103.4153
24.2134 | -103.4192
24.2171 | -103.4249
24.2208 | -103.4325
24.2245 | -103.4395
24.2283 | -103.4432
24.232 | -103.4454
24.2357 | -103.4475
24.2394 | -103.4494
24.2431 | -103.4513
24.2468 | -103.4532
24.2505 | -103.4549
24.2542 | -103.4542
24.2579 | -103.4478
24.2617 | -103.4156
24.2654 | -103.3618
24.2691 | -103.3021
24.2728 | -103.2132
24.2765 | -103.1
24.2802 | -102.9632
24.2839 | -102.8411
24.2876 | -102.727
24.2913 | -102.6071
24.2951 | -102.4916
24.2988 | -102.3701
24.3025 | -102.2525
24.3062 | -102.1364
24.3099 | -102.0182
24.3136 | -101.8947
24.3173 | -101.7656
24.321 | -101.6332
24.3248 | -101.4956
24.3285 | -101.3539
24.3322 | -101.2096
24.3359 | -101.0636
24.3396 | -100.9165
24.3433 | -100.767
24.347 | -100.6164
24.3507 | -100.4658
24.3544 | -100.3152
24.3582 | -100.165
24.3619 | -100.0178
24.3656 | -99.8729
24.3693 | -99.7304
24.373 | -99.5905
24.3767 | -99.4537
24.3804 | -99.3208
24.3841 | -99.1922
24.3878 | -99.0687
24.3916 | -98.9498
24.3953 | -98.8363
24.399 | -98.7292
24.4027 | -98.6286
24.4064 | -98.5337
24.4101 | -98.4452
24.4138 | -98.3638
24.4175 | -98.2893
24.4212 | -98.2214
24.425 | -98.1612
24.4287 | -98.1092
24.4324 | -98.0647
24.4361 | -98.0274
24.4398 | -97.9978
24.4435 | -97.9758
24.4472 | -97.9616
24.4509 | -97.9555
24.4547 | -97.958
24.4584 | -97.9682
24.4621 | -97.9856
24.4658 | -98.0109
24.4695 | -98.0436
24.4732 | -98.0838
24.4769 | -98.1316
24.4806 | -98.1868
24.4843 | -98.2489
24.4881 | -98.3185
24.4918 | -98.3948
24.4955 | -98.4776
24.4992 | -98.5676
24.5029 | -98.6634
24.5066 | -98.7651
24.5103 | -98.8722
24.514 | -98.9853
24.5177 | -99.1036
24.5215 | -99.2266
24.5252 | -99.3539
24.5289 | -99.4843
24.5326 | -99.6181
24.5363 | -99.7552
24.54 | -99.8946
24.5437 | -100.0355
24.5474 | -100.1759
24.5511 | -100.3132
24.5549 | -100.4474
24.5586 | -100.571
24.5623 | -100.6849
24.566 | -100.7893
24.5697 | -100.8867
24.5734 | -100.9726
24.5771 | -101.0426
24.5808 | -101.096
24.5846 | -101.1509
24.5883 | -101.2067
24.592 | -101.2544
24.5957 | -101.2934
24.5994 | -101.3235
24.6031 | -101.3458
24.6068 | -101.3692
24.6105 | -101.388
24.6142 | -101.3899
24.618 | -101.3892
24.6217 | -101.3816
24.6254 | -101.3667
24.6291 | -101.2176
24.6328 | -101.1709
24.6365 | -101.1028
24.6402 | -101.0245
24.6439 | -100.9318
24.6476 | -100.8348
24.6514 | -100.7271
24.6551 | -100.6123
24.6588 | -100.496
24.6625 | -100.3812
24.6662 | -100.2671
24.6699 | -100.1542
24.6736 | -100.0437
24.6773 | -99.936
24.681 | -99.8312
24.6848 | -99.7299
24.6885 | -99.6326
24.6922 | -99.54
24.6959 | -99.4516
24.6996 | -99.3684
24.7033 | -99.2907
24.707 | -99.2174
24.7107 | -99.149
24.7145 | -99.0861
24.7182 | -99.0294
24.7219 | -98.9797
24.7256 | -98.9359
24.7293 | -98.8987
24.733 | -98.8683
24.7367 | -98.8439
24.7404 | -98.826
24.7441 | -98.8145
24.7479 | -98.8093
24.7516 | -98.8101
24.7553 | -98.8183
24.759 | -98.8329
24.7627 | -98.8534
24.7664 | -98.8806
24.7701 | -98.9149
24.7738 | -98.9548
24.7775 | -99.0002
24.7813 | -99.0521
24.785 | -99.1097
24.7887 | -99.1729
24.7924 | -99.2413
24.7961 | -99.3156
24.7998 | -99.3949
24.8035 | -99.479
24.8072 | -99.5672
24.8109 | -99.6597
24.8147 | -99.7559
24.8184 | -99.8571
24.8221 | -99.9624
24.8258 | -100.0705
24.8295 | -100.1816
24.8332 | -100.2946
24.8369 | -100.4083
24.8406 | -100.524
24.8444 | -100.6405
24.8481 | -100.7576
24.8518 | -100.8761
24.8555 | -100.9949
24.8592 | -101.1132
24.8629 | -101.2302
24.8666 | -101.3461
24.8703 | -101.4599
24.874 | -101.5729
24.8778 | -101.6839
24.8815 | -101.7924
24.8852 | -101.8973
24.8889 | -102.0004
24.8926 | -102.1011
24.8963 | -102.1995
24.9 | -102.2983
24.9037 | -102.3976
24.9074 | -102.5041
24.9112 | -102.6103
24.9149 | -102.7134
24.9186 | -102.8093
24.9223 | -102.9045
24.926 | -103.001
24.9297 | -103.0946
24.9334 | -103.1579
24.9371 | -103.1968
24.9408 | -103.2284
24.9446 | -103.2536
24.9483 | -103.2745
24.952 | -103.2915
24.9557 | -103.3093
24.9594 | -103.3259
24.9631 | -103.3441
24.9668 | -103.3624
24.9705 | -103.3767
24.9743 | -103.385
24.978 | -103.3908
24.9817 | -103.3916
24.9854 | -103.3914
24.9891 | -103.3909
24.9928 | -103.3902
24.9965 | -103.388
25.0002 | -103.3838
25.0039 | -103.3803
25.0077 | -103.3755
25.0114 | -103.3684
25.0151 | -103.3588
25.0188 | -103.3521
25.0225 | -103.347
25.0262 | -103.3399
25.0299 | -103.3328
25.0336 | -103.3264
25.0373 | -103.318
25.0411 | -103.3091
25.0448 | -103.2989
25.0485 | -103.2881
25.0522 | -103.2764
25.0559 | -103.2666
25.0596 | -103.2581
25.0633 | -103.249
25.067 | -103.24
25.0707 | -103.2305
25.0745 | -103.2202
25.0782 | -103.2082
25.0819 | -103.198
25.0856 | -103.1886
25.0893 | -103.1803
25.093 | -103.1722
25.0967 | -103.1651
25.1004 | -103.1601
25.1042 | -103.1544
25.1079 | -103.1482
25.1116 | -103.1434
25.1153 | -103.1404
25.119 | -103.1383
25.1227 | -103.136
25.1264 | -103.135
25.1301 | -103.1351
25.1338 | -103.1351
25.1376 | -103.1366
25.1413 | -103.1381
25.145 | -103.1398
25.1487 | -103.1421
25.1524 | -103.1449
25.1561 | -103.1467
25.1598 | -103.1512
25.1635 | -103.1613
25.1672 | -103.1666
25.171 | -103.1722
25.1747 | -103.177
25.1784 | -103.1836
25.1821 | -103.189
25.1858 | -103.1937
25.1895 | -103.2008
25.1932 | -103.2089
25.1969 | -103.2176
25.2006 | -103.2267
25.2044 | -103.2326
25.2081 | -103.2395
25.2118 | -103.2475
25.2155 | -103.2555
25.2192 | -103.2628
25.2229 | -103.2695
25.2266 | -103.278
25.2303 | -103.2859
25.2341 | -103.293
25.2378 | -103.2991
25.2415 | -103.3048
25.2452 | -103.3098
25.2489 | -103.3128
25.2526 | -103.3172
25.2563 | -103.3219
25.26 | -103.3255
25.2637 | -103.3289
25.2675 | -103.3329
25.2712 | -103.3363
25.2749 | -103.3383
25.2786 | -103.3389
25.2823 | -103.3365
25.286 | -103.3338
25.2897 | -103.3323
25.2934 | -103.3306
25.2971 | -103.3274
25.3009 | -103.322
25.3046 | -103.3159
25.3083 | -103.3104
25.312 | -103.3041
25.3157 | -103.2959
25.3194 | -103.2885
25.3231 | -103.2817
25.3268 | -103.277
25.3305 | -103.273
25.3343 | -103.265
25.338 | -103.2551
25.3417 | -103.2437
25.3454 | -103.2311
25.3491 | -103.2174
25.3528 | -103.2035
25.3565 | -103.1891
25.3602 | -103.175
25.364 | -103.1624
25.3677 | -103.1502
25.3714 | -103.137
25.3751 | -103.1229
25.3788 | -103.108
25.3825 | -103.0939
25.3862 | -103.0804
25.3899 | -103.0655
25.3936 | -103.0516
25.3974 | -103.0373
25.4011 | -103.0242
25.4048 | -103.0126
25.4085 | -103.0011
25.4122 | -102.9881
25.4159 | -102.9765
25.4196 | -102.965
25.4233 | -102.9545
25.427 | -102.9442
25.4308 | -102.9339
25.4345 | -102.9258
25.4382 | -102.9168
25.4419 | -102.9094
25.4456 | -102.9012
25.4493 | -102.8928
25.453 | -102.8858
25.4567 | -102.8798
25.4605 | -102.8743
25.4642 | -102.8692
25.4679 | -102.8649
25.4716 | -102.8632
25.4753 | -102.8618
25.479 | -102.8597
25.4827 | -102.8597
25.4864 | -102.8595
25.4901 | -102.8587
25.4939 | -102.8592
25.4976 | -102.8616
25.5013 | -102.866
25.505 | -102.87
25.5087 | -102.8742
25.5124 | -102.88
25.5161 | -102.8848
25.5198 | -102.8907
25.5235 | -102.8979
25.5273 | -102.9053
25.531 | -102.911
25.5347 | -102.9177
25.5384 | -102.9247
25.5421 | -102.933
25.5458 | -102.9409
25.5495 | -102.9473
25.5532 | -102.9545
25.5569 | -102.961
25.5607 | -102.9673
25.5644 | -102.9765
25.5681 | -102.9838
25.5718 | -102.9872
25.5755 | -102.99
25.5792 | -102.9956
25.5829 | -103
25.5866 | -103.0033
25.5904 | -103.0062
25.5941 | -103.0094
25.5978 | -103.0123
25.6015 | -103.0141
25.6052 | -103.016
25.6089 | -103.0179
25.6126 | -103.0192
25.6163 | -103.02
25.62 | -103.0182
25.6238 | -103.0131
25.6275 | -103.0064
25.6312 | -103.0014
25.6349 | -102.9969
25.6386 | -102.9899
25.6423 | -102.9856
25.646 | -102.9799
25.6497 | -102.9741
25.6534 | -102.9678
25.6572 | -102.9596
25.6609 | -102.9526
25.6646 | -102.9446
25.6683 | -102.9351
25.672 | -102.9248
25.6757 | -102.9164
25.6794 | -102.9057
25.6831 | -102.8938
25.6868 | -102.8829
25.6906 | -102.8733
25.6943 | -102.8654
25.698 | -102.8571
25.7017 | -102.8466
25.7054 | -102.8386
25.7091 | -102.8307
25.7128 | -102.8227
25.7165 | -102.8168
25.7203 | -102.8114
25.724 | -102.8037
25.7277 | -102.7945
25.7314 | -102.7863
25.7351 | -102.7811
25.7388 | -102.7759
25.7425 | -102.7706
25.7462 | -102.7647
25.7499 | -102.7589
25.7537 | -102.7547
25.7574 | -102.7509
25.7611 | -102.7476
25.7648 | -102.7457
25.7685 | -102.7429
25.7722 | -102.7414
25.7759 | -102.74
25.7796 | -102.7393
25.7833 | -102.7398
25.7871 | -102.7399
25.7908 | -102.7408
25.7945 | -102.7421
25.7982 | -102.7443
25.8019 | -102.7457
25.8056 | -102.7456
25.8093 | -102.7497
25.813 | -102.7571
25.8167 | -102.7633
25.8205 | -102.7682
25.8242 | -102.7739
25.8279 | -102.7799
25.8316 | -102.7863
25.8353 | -102.7906
25.839 | -102.7954
25.8427 | -102.8014
25.8464 | -102.8054
25.8502 | -102.8123
25.8539 | -102.8202
25.8576 | -102.8268
25.8613 | -102.8337
25.865 | -102.8402
25.8687 | -102.8464
25.8724 | -102.8524
25.8761 | -102.8587
25.8798 | -102.8652
25.8836 | -102.8716
25.8873 | -102.8757
25.891 | -102.88
25.8947 | -102.8878
25.8984 | -102.8948
25.9021 | -102.9008
25.9058 | -102.9062
25.9095 | -102.9128
25.9132 | -102.9203
25.917 | -102.9265
25.9207 | -102.9327
25.9244 | -102.938
25.9281 | -102.9439
25.9318 | -102.9512
25.9355 | -102.9594
25.9392 | -102.9666
25.9429 | -102.9718
25.9466 | -102.9786
25.9504 | -102.9853
25.9541 | -102.9925
25.9578 | -102.9994
25.9615 | -103.0079
25.9652 | -103.0153
25.9689 | -103.0217
25.9726 | -103.0305
25.9763 | -103.0389
25.9801 | -103.0478
25.9838 | -103.0573
25.9875 | -103.0684
25.9912 | -103.0802
25.9949 | -103.0919
25.9986 | -103.1014
26.0023 | -103.1119
26.006 | -103.1204
26.0097 | -103.1308
26.0135 | -103.1399
26.0172 | -103.1499
26.0209 | -103.1632
26.0246 | -103.175
26.0283 | -103.1877
26.032 | -103.2001
26.0357 | -103.2129
26.0394 | -103.2258
26.0431 | -103.2397
26.0469 | -103.2518
26.0506 | -103.2621
26.0543 | -103.2721
26.058 | -103.2814
26.0617 | -103.2912
26.0654 | -103.3036
26.0691 | -103.3171
26.0728 | -103.3258
26.0765 | -103.3344
26.0803 | -103.3443
26.084 | -103.35
26.0877 | -103.355
26.0914 | -103.3631
26.0951 | -103.3715
26.0988 | -103.3787
26.1025 | -103.3838
26.1062 | -103.3873
26.11 | -103.3908
26.1137 | -103.3965
26.1174 | -103.4
26.1211 | -103.4022
26.1248 | -103.4028
26.1285 | -103.4046
26.1322 | -103.4032
26.1359 | -103.4019
26.1396 | -103.4
26.1434 | -103.3981
26.1471 | -103.3942
26.1508 | -103.3904
26.1545 | -103.3875
26.1582 | -103.3839
26.1619 | -103.3793
26.1656 | -103.3738
26.1693 | -103.3678
26.173 | -103.3626
26.1768 | -103.3541
26.1805 | -103.3448
26.1842 | -103.3354
26.1879 | -103.3254
26.1916 | -103.3169
26.1953 | -103.3097
26.199 | -103.3018
26.2027 | -103.2935
26.2064 | -103.2846
26.2102 | -103.277
26.2139 | -103.27
26.2176 | -103.262
26.2213 | -103.2535
26.225 | -103.2453
26.2287 | -103.2377
26.2324 | -103.2301
26.2361 | -103.2252
26.2399 | -103.2176
26.2436 | -103.2093
26.2473 | -103.2009
26.251 | -103.1932
26.2547 | -103.1854
26.2584 | -103.1815
26.2621 | -103.1783
26.2658 | -103.1732
26.2695 | -103.1688
26.2733 | -103.1649
26.277 | -103.1608
26.2807 | -103.1557
26.2844 | -103.1517
26.2881 | -103.1496
26.2918 | -103.1457
26.2955 | -103.1418
26.2992 | -103.1381
26.3029 | -103.1351
26.3067 | -103.134
26.3104 | -103.135
26.3141 | -103.1355
26.3178 | -103.135
26.3215 | -103.1351
26.3252 | -103.1353
26.3289 | -103.1365
26.3326 | -103.1377
26.3363 | -103.1387
26.3401 | -103.1391
26.3438 | -103.1378
26.3475 | -103.1369
26.3512 | -103.138
26.3549 | -103.1396
26.3586 | -103.1414
26.3623 | -103.1435
26.366 | -103.1444
26.3698 | -103.1454
26.3735 | -103.1473
26.3772 | -103.1518
26.3809 | -103.1545
26.3846 | -103.1557
26.3883 | -103.1577
26.392 | -103.1582
26.3957 | -103.1593
26.3994 | -103.1595
26.4032 | -103.1595
26.4069 | -103.1602
26.4106 | -103.16
26.4143 | -103.1602
26.418 | -103.1618
26.4217 | -103.1633
26.4254 | -103.1644
26.4291 | -103.1652
26.4328 | -103.1653
26.4366 | -103.1664
26.4403 | -103.1678
26.444 | -103.1697
26.4477 | -103.1717
26.4514 | -103.1725
26.4551 | -103.1758
26.4588 | -103.1783
26.4625 | -103.1786
26.4662 | -103.1789
26.47 | -103.1796
26.4737 | -103.1809
26.4774 | -103.183
26.4811 | -103.1857
26.4848 | -103.1864
26.4885 | -103.1855
26.4922 | -103.1855
26.4959 | -103.188
26.4997 | -103.1906
26.5034 | -103.1935
26.5071 | -103.1957
26.5108 | -103.1985
26.5145 | -103.2015
26.5182 | -103.2052
26.5219 | -103.2076
26.5256 | -103.2086
26.5293 | -103.211
26.5331 | -103.2134
26.5368 | -103.2172
26.5405 | -103.2224
26.5442 | -103.226
26.5479 | -103.228
26.5516 | -103.2298
26.5553 | -103.2336
26.559 | -103.2392
26.5627 | -103.2448
26.5665 | -103.2488
26.5702 | -103.2538
26.5739 | -103.2595
26.5776 | -103.2658
26.5813 | -103.2715
26.585 | -103.2778
26.5887 | -103.2831
26.5924 | -103.2878
26.5961 | -103.2936
26.5999 | -103.2992
26.6036 | -103.3042
26.6073 | -103.3123
26.611 | -103.3176
26.6147 | -103.3229
26.6184 | -103.3284
26.6221 | -103.3342
26.6258 | -103.3396
26.6296 | -103.3448
26.6333 | -103.3499
26.637 | -103.3542
26.6407 | -103.3558
26.6444 | -103.3581
26.6481 | -103.3614
26.6518 | -103.3646
26.6555 | -103.3684
26.6592 | -103.3719
26.663 | -103.3743
26.6667 | -103.3769
26.6704 | -103.3806
26.6741 | -103.3818
26.6778 | -103.3823
26.6815 | -103.3828
26.6852 | -103.3823
26.6889 | -103.383
26.6926 | -103.3832
26.6964 | -103.3835
26.7001 | -103.3834
26.7038 | -103.3824
26.7075 | -103.3806
26.7112 | -103.3813
26.7149 | -103.3827
26.7186 | -103.3841
26.7223 | -103.3838
26.7261 | -103.3819
26.7298 | -103.3809
26.7335 | -103.3777
26.7372 | -103.3755
26.7409 | -103.3743
26.7446 | -103.3742
26.7483 | -103.3733
26.752 | -103.3692
26.7557 | -103.3681
26.7595 | -103.3664
26.7632 | -103.3641
26.7669 | -103.3628
26.7706 | -103.3643
26.7743 | -103.3641
26.778 | -103.3632
26.7817 | -103.364
26.7854 | -103.364
26.7891 | -103.3655
26.7929 | -103.3665
26.7966 | -103.3677
26.8003 | -103.3699
26.804 | -103.3717
26.8077 | -103.3745
26.8114 | -103.3775
26.8151 | -103.3794
26.8188 | -103.3813
26.8225 | -103.3843
26.8263 | -103.3879
26.83 | -103.3917
26.8337 | -103.3948
26.8374 | -103.3982
26.8411 | -103.4033
26.8448 | -103.4096
26.8485 | -103.4149
26.8522 | -103.4194
26.8559 | -103.4239
26.8597 | -103.4277
26.8634 | -103.4336
26.8671 | -103.4426
26.8708 | -103.4495
26.8745 | -103.4552
26.8782 | -103.4609
26.8819 | -103.4668
26.8856 | -103.4741
26.8894 | -103.4795
26.8931 | -103.4832
26.8968 | -103.488
26.9005 | -103.4952
26.9042 | -103.504
26.9079 | -103.5106
26.9116 | -103.5162
26.9153 | -103.5214
26.919 | -103.5256
26.9228 | -103.5313
26.9265 | -103.5389
26.9302 | -103.5449
26.9339 | -103.5491
26.9376 | -103.5526
26.9413 | -103.5559
26.945 | -103.5591
26.9487 | -103.5612
26.9524 | -103.5644
26.9562 | -103.5681
26.9599 | -103.5707
26.9636 | -103.5729
26.9673 | -103.576
26.971 | -103.5799
26.9747 | -103.5835
26.9784 | -103.5856
26.9821 | -103.5866
26.9858 | -103.5858
26.9896 | -103.5854
26.9933 | -103.587
26.997 | -103.5886
27.0007 | -103.5888
27.0044 | -103.5889
27.0081 | -103.589
27.0118 | -103.5875
27.0155 | -103.5849
27.0193 | -103.5856
27.023 | -103.5843
27.0267 | -103.5814
27.0304 | -103.58
27.0341 | -103.5802
27.0378 | -103.5777
27.0415 | -103.5757
27.0452 | -103.5752
27.0489 | -103.5725
27.0527 | -103.5716
27.0564 | -103.5699
27.0601 | -103.5684
27.0638 | -103.5672
27.0675 | -103.565
27.0712 | -103.5609
27.0749 | -103.5568
27.0786 | -103.5549
27.0823 | -103.5524
27.0861 | -103.5505
27.0898 | -103.5486
27.0935 | -103.5466
27.0972 | -103.5448
27.1009 | -103.5435
27.1046 | -103.5433
27.1083 | -103.5432
27.112 | -103.5439
27.1157 | -103.5425
27.1195 | -103.5408
27.1232 | -103.5399
27.1269 | -103.5386
27.1306 | -103.5391
27.1343 | -103.536
27.138 | -103.5341
27.1417 | -103.5325
27.1454 | -103.5326
27.1492 | -103.5316
27.1529 | -103.5297
27.1566 | -103.5287
27.1603 | -103.5291
27.164 | -103.5281
27.1677 | -103.527
27.1714 | -103.5249
27.1751 | -103.5215
27.1788 | -103.5186
27.1826 | -103.517
27.1863 | -103.5167
27.19 | -103.5161
27.1937 | -103.5162
27.1974 | -103.5163
27.2011 | -103.5155
27.2048 | -103.5145
27.2085 | -103.5125
27.2122 | -103.5117
27.216 | -103.5107
27.2197 | -103.5116
27.2234 | -103.5109
27.2271 | -103.5088
27.2308 | -103.5072
27.2345 | -103.5069
27.2382 | -103.5077
27.2419 | -103.5086
27.2457 | -103.5078
27.2494 | -103.5056
27.2531 | -103.5049
27.2568 | -103.5045
27.2605 | -103.5044
27.2642 | -103.504
27.2679 | -103.502
27.2716 | -103.5016
27.2753 | -103.5012
27.2791 | -103.5007
27.2828 | -103.4991
27.2865 | -103.4962
27.2902 | -103.495
27.2939 | -103.4947
27.2976 | -103.4938
27.3013 | -103.4931
27.305 | -103.4911
27.3087 | -103.4874
27.3125 | -103.485
27.3162 | -103.484
27.3199 | -103.4832
27.3236 | -103.4827
27.3273 | -103.481
27.331 | -103.4798
27.3347 | -103.4782
27.3384 | -103.4759
27.3421 | -103.4727
27.3459 | -103.4703
27.3496 | -103.4677
27.3533 | -103.4659
27.357 | -103.4635
27.3607 | -103.461
27.3644 | -103.4587
27.3681 | -103.4556
27.3718 | -103.4513
27.3756 | -103.4488
27.3793 | -103.447
27.383 | -103.4455
27.3867 | -103.4424
27.3904 | -103.4381
27.3941 | -103.4343
27.3978 | -103.4314
27.4015 | -103.4294
27.4052 | -103.4267
27.409 | -103.4235
27.4127 | -103.4209
27.4164 | -103.418
27.4201 | -103.4129
27.4238 | -103.4092
27.4275 | -103.4059
27.4312 | -103.4012
27.4349 | -103.3982
27.4386 | -103.3945
27.4424 | -103.3897
27.4461 | -103.3848
27.4498 | -103.3795
27.4535 | -103.375
27.4572 | -103.3736
27.4609 | -103.3699
27.4646 | -103.3641
27.4683 | -103.3604
27.472 | -103.356
27.4758 | -103.3511
27.4795 | -103.3458
27.4832 | -103.3413
27.4869 | -103.3372
27.4906 | -103.3312
27.4943 | -103.3256
27.498 | -103.3194
27.5017 | -103.3142
27.5055 | -103.3083
27.5092 | -103.301
27.5129 | -103.2966
27.5166 | -103.2921
27.5203 | -103.2884
27.524 | -103.284
27.5277 | -103.2787
27.5314 | -103.2728
27.5351 | -103.269
27.5389 | -103.2659
27.5426 | -103.2626
27.5463 | -103.2583
27.55 | -103.2537
27.5537 | -103.2488
27.5574 | -103.2444
27.5611 | -103.2398
27.5648 | -103.2357
27.5685 | -103.231
27.5723 | -103.2254
27.576 | -103.2195
27.5797 | -103.2142
27.5834 | -103.2099
27.5871 | -103.2061
27.5908 | -103.2033
27.5945 | -103.2023
27.5982 | -103.2
27.6019 | -103.197
27.6057 | -103.1925
27.6094 | -103.1892
27.6131 | -103.1868
27.6168 | -103.1857
27.6205 | -103.1842
27.6242 | -103.1817
27.6279 | -103.1785
27.6316 | -103.1755
27.6354 | -103.1743
27.6391 | -103.1716
27.6428 | -103.1674
27.6465 | -103.1641
27.6502 | -103.1624
27.6539 | -103.1621
27.6576 | -103.1619
27.6613 | -103.1626
27.665 | -103.1626
27.6688 | -103.161
27.6725 | -103.1604
27.6762 | -103.1602
27.6799 | -103.1597
27.6836 | -103.1583
27.6873 | -103.1587
27.691 | -103.1587
27.6947 | -103.1593
27.6984 | -103.1605
27.7022 | -103.1608
27.7059 | -103.1612
27.7096 | -103.1609
27.7133 | -103.1617
27.717 | -103.1639
27.7207 | -103.1645
27.7244 | -103.1654
27.7281 | -103.1667
27.7318 | -103.1689
27.7356 | -103.1708
27.7393 | -103.1741
27.743 | -103.1763
27.7467 | -103.1791
27.7504 | -103.1821
27.7541 | -103.1863
27.7578 | -103.1924
27.7615 | -103.1981
27.7653 | -103.2004
27.769 | -103.202
27.7727 | -103.2045
27.7764 | -103.2101
27.7801 | -103.2158
27.7838 | -103.2189
27.7875 | -103.2215
27.7912 | -103.2249
27.7949 | -103.228
27.7987 | -103.2334
27.8024 | -103.2387
27.8061 | -103.2437
27.8098 | -103.2484
27.8135 | -103.2534
27.8172 | -103.2585
27.8209 | -103.2645
27.8246 | -103.2696
27.8283 | -103.2733
27.8321 | -103.2796
27.8358 | -103.2855
27.8395 | -103.2877
27.8432 | -103.2942
27.8469 | -103.3008
27.8506 | -103.3058
27.8543 | -103.3093
27.858 | -103.307
27.8617 | -103.3063
27.8655 | -103.3101
27.8692 | -103.3138
27.8729 | -103.3166
27.8766 | -103.3198
27.8803 | -103.3225
27.884 | -103.3252
27.8877 | -103.3276
27.8914 | -103.3315
27.8952 | -103.3335
27.8989 | -103.3351
27.9026 | -103.3381
27.9063 | -103.3388
27.91 | -103.3388
27.9137 | -103.3401
27.9174 | -103.341
27.9211 | -103.3393
27.9248 | -103.3329
27.9286 | -103.3216
27.9323 | -103.312
27.936 | -103.3079
27.9397 | -103.3025
27.9434 | -103.2985
27.9471 | -103.2961
27.9508 | -103.2921
27.9545 | -103.2867
27.9582 | -103.2796
27.962 | -103.2719
27.9657 | -103.2647
27.9694 | -103.2584
27.9731 | -103.251
27.9768 | -103.2425
27.9805 | -103.2334
27.9842 | -103.2214
27.9879 | -103.2104
27.9916 | -103.2001
27.9954 | -103.1885
27.9991 | -103.1744
28.0028 | -103.1609