## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 261 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 251
Hood | 02-92B/P | 211
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1005.2 | hPa |  | PRC PRO 01
Humidity | 60.1 | %/Hr
Internal temperature of calibration bench | 23.44 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1011 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.6 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.12 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.99 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3104 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 04/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 09/07/2018
Technical Manager | Anthony HUE | Approval |  | 16/07/2017
CEO | Robert GLEMOT | Approval |  | 16/07/2017
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3882 | 57.378
0.012209 | -35.9395 | 38.3978
0.018313 | -34.8856 | 28.449
0.024417 | -34.438 | 22.0574
0.030521 | -34.335 | 18.2854
0.036626 | -34.194 | 15.1976
0.04273 | -34.1109 | 12.8914
0.048834 | -34.0588 | 11.3099
0.054938 | -33.9888 | 10.1088
0.061043 | -33.9132 | 8.9802
0.067147 | -33.9019 | 8.1088
0.073251 | -33.9028 | 7.3567
0.079355 | -33.9208 | 6.8012
0.08546 | -33.9388 | 6.3454
0.091564 | -33.9354 | 5.9718
0.097668 | -33.9013 | 5.7176
0.10377 | -33.8532 | 5.3721
0.10988 | -33.822 | 5.009
0.11598 | -33.8127 | 4.647
0.12209 | -33.8295 | 4.3659
0.12819 | -33.8623 | 4.1327
0.13429 | -33.8942 | 3.8939
0.1404 | -33.9116 | 3.6245
0.1465 | -33.9276 | 3.3693
0.15261 | -33.9352 | 3.1323
0.15871 | -33.9323 | 2.937
0.16482 | -33.9182 | 2.7777
0.17092 | -33.8938 | 2.7134
0.17702 | -33.875 | 2.6147
0.18313 | -33.8578 | 2.5311
0.18923 | -33.8453 | 2.4251
0.19534 | -33.8384 | 2.3229
0.20144 | -33.8373 | 2.2118
0.20754 | -33.8399 | 2.1409
0.21365 | -33.8421 | 2.0683
0.21975 | -33.8462 | 2.0056
0.22586 | -33.854 | 1.942
0.23196 | -33.8623 | 1.8696
0.23807 | -33.8708 | 1.7913
0.24417 | -33.8728 | 1.6957
0.25027 | -33.8759 | 1.6099
0.25638 | -33.8757 | 1.5163
0.26248 | -33.8737 | 1.4282
0.26859 | -33.8713 | 1.3442
0.27469 | -33.8707 | 1.2578
0.2808 | -33.8768 | 1.195
0.2869 | -33.8815 | 1.1277
0.293 | -33.8872 | 1.0717
0.29911 | -33.8937 | 1.0345
0.30521 | -33.8999 | 0.99921
0.31132 | -33.9038 | 0.98663
0.31742 | -33.9081 | 0.96568
0.32353 | -33.9114 | 0.94878
0.32963 | -33.9141 | 0.93467
0.33573 | -33.9157 | 0.92239
0.34184 | -33.918 | 0.90674
0.34794 | -33.921 | 0.87193
0.35405 | -33.9241 | 0.84221
0.36015 | -33.9275 | 0.80905
0.36626 | -33.9297 | 0.7722
0.37236 | -33.9321 | 0.72255
0.37846 | -33.9343 | 0.66564
0.38457 | -33.9334 | 0.60713
0.39067 | -33.934 | 0.55739
0.39678 | -33.9324 | 0.50159
0.40288 | -33.9301 | 0.449
0.40899 | -33.9263 | 0.39772
0.41509 | -33.9219 | 0.34605
0.42119 | -33.9168 | 0.29901
0.4273 | -33.9131 | 0.25043
0.4334 | -33.9096 | 0.2009
0.43951 | -33.9065 | 0.15217
0.44561 | -33.9049 | 0.10645
0.45172 | -33.9064 | 0.068257
0.45782 | -33.9061 | 0.026037
0.46392 | -33.9072 | -0.0128
0.47003 | -33.9089 | -0.048253
0.47613 | -33.9109 | -0.080614
0.48224 | -33.913 | -0.11138
0.48834 | -33.9155 | -0.13287
0.49445 | -33.9175 | -0.15759
0.50055 | -33.9201 | -0.17992
0.50665 | -33.9218 | -0.19789
0.51276 | -33.923 | -0.21731
0.51886 | -33.9242 | -0.23388
0.52497 | -33.9233 | -0.23944
0.53107 | -33.9237 | -0.2489
0.53717 | -33.9241 | -0.25933
0.54328 | -33.9239 | -0.26357
0.54938 | -33.9239 | -0.27087
0.55549 | -33.9234 | -0.27904
0.56159 | -33.9233 | -0.28981
0.5677 | -33.9229 | -0.29972
0.5738 | -33.9224 | -0.31253
0.5799 | -33.9224 | -0.32432
0.58601 | -33.9228 | -0.34083
0.59211 | -33.9238 | -0.36393
0.59822 | -33.9245 | -0.38095
0.60432 | -33.9255 | -0.40282
0.61043 | -33.9265 | -0.42456
0.61653 | -33.9277 | -0.4466
0.62263 | -33.9288 | -0.47027
0.62874 | -33.9303 | -0.49113
0.63484 | -33.9314 | -0.51274
0.64095 | -33.9327 | -0.53427
0.64705 | -33.9339 | -0.55477
0.65316 | -33.935 | -0.57479
0.65926 | -33.9361 | -0.59762
0.66536 | -33.9356 | -0.61602
0.67147 | -33.9359 | -0.63258
0.67757 | -33.9365 | -0.6507
0.68368 | -33.9363 | -0.67069
0.68978 | -33.9364 | -0.68866
0.69589 | -33.9361 | -0.70682
0.70199 | -33.9353 | -0.72828
0.70809 | -33.9347 | -0.74821
0.7142 | -33.9342 | -0.76717
0.7203 | -33.9334 | -0.78646
0.72641 | -33.9326 | -0.8069
0.73251 | -33.9323 | -0.82733
0.73862 | -33.9323 | -0.85118
0.74472 | -33.9326 | -0.87008
0.75082 | -33.9321 | -0.89067
0.75693 | -33.9322 | -0.91366
0.76303 | -33.9321 | -0.93563
0.76914 | -33.9325 | -0.96233
0.77524 | -33.9327 | -0.98695
0.78135 | -33.9331 | -1.01
0.78745 | -33.9336 | -1.0346
0.79355 | -33.9341 | -1.0589
0.79966 | -33.935 | -1.0845
0.80576 | -33.9361 | -1.1149
0.81187 | -33.9363 | -1.1442
0.81797 | -33.9369 | -1.1748
0.82408 | -33.9375 | -1.2019
0.83018 | -33.938 | -1.2331
0.83628 | -33.9388 | -1.2636
0.84239 | -33.9386 | -1.2918
0.84849 | -33.9385 | -1.3198
0.8546 | -33.9389 | -1.3475
0.8607 | -33.9389 | -1.3769
0.86681 | -33.9389 | -1.4053
0.87291 | -33.9384 | -1.4311
0.87901 | -33.9378 | -1.4534
0.88512 | -33.9376 | -1.4765
0.89122 | -33.9374 | -1.4999
0.89733 | -33.9371 | -1.5192
0.90343 | -33.9368 | -1.5405
0.90953 | -33.9369 | -1.5553
0.91564 | -33.9363 | -1.5711
0.92174 | -33.9359 | -1.5853
0.92785 | -33.9354 | -1.5986
0.93395 | -33.9353 | -1.6102
0.94006 | -33.9353 | -1.6236
0.94616 | -33.9351 | -1.6279
0.95226 | -33.9351 | -1.6391
0.95837 | -33.9352 | -1.6506
0.96447 | -33.9356 | -1.6605
0.97058 | -33.9356 | -1.6701
0.97668 | -33.9353 | -1.6776
0.98279 | -33.9354 | -1.6807
0.98889 | -33.9356 | -1.6851
0.99499 | -33.9356 | -1.6933
1.0011 | -33.9359 | -1.6977
1.0072 | -33.9358 | -1.701
1.0133 | -33.9357 | -1.7045
1.0194 | -33.9357 | -1.7095
1.0255 | -33.9354 | -1.716
1.0316 | -33.9348 | -1.7197
1.0377 | -33.9343 | -1.7252
1.0438 | -33.9336 | -1.7292
1.0499 | -33.9321 | -1.7352
1.056 | -33.9309 | -1.7396
1.0621 | -33.9297 | -1.7449
1.0682 | -33.9293 | -1.7495
1.0743 | -33.9281 | -1.757
1.0805 | -33.9275 | -1.7635
1.0866 | -33.9259 | -1.774
1.0927 | -33.9245 | -1.7812
1.0988 | -33.9237 | -1.7915
1.1049 | -33.9223 | -1.7987
1.111 | -33.9207 | -1.8095
1.1171 | -33.9193 | -1.8186
1.1232 | -33.918 | -1.8327
1.1293 | -33.9168 | -1.8443
1.1354 | -33.9158 | -1.8565
1.1415 | -33.9146 | -1.8703
1.1476 | -33.9133 | -1.8841
1.1537 | -33.9119 | -1.8972
1.1598 | -33.9118 | -1.9154
1.1659 | -33.9113 | -1.9272
1.172 | -33.9105 | -1.9403
1.1781 | -33.9098 | -1.9542
1.1842 | -33.9092 | -1.9697
1.1903 | -33.91 | -1.9849
1.1964 | -33.9099 | -1.9993
1.2025 | -33.9098 | -2.0144
1.2086 | -33.9103 | -2.03
1.2147 | -33.9106 | -2.045
1.2209 | -33.9113 | -2.0584
1.227 | -33.9132 | -2.0734
1.2331 | -33.9142 | -2.0876
1.2392 | -33.9152 | -2.1034
1.2453 | -33.9164 | -2.1153
1.2514 | -33.9178 | -2.1275
1.2575 | -33.9194 | -2.1415
1.2636 | -33.9217 | -2.153
1.2697 | -33.9235 | -2.1646
1.2758 | -33.9251 | -2.1787
1.2819 | -33.9269 | -2.1922
1.288 | -33.9286 | -2.206
1.2941 | -33.9307 | -2.2167
1.3002 | -33.9327 | -2.2249
1.3063 | -33.9344 | -2.2349
1.3124 | -33.9363 | -2.2449
1.3185 | -33.9383 | -2.2556
1.3246 | -33.9401 | -2.2663
1.3307 | -33.9423 | -2.2758
1.3368 | -33.9441 | -2.2833
1.3429 | -33.9458 | -2.292
1.349 | -33.9473 | -2.3005
1.3551 | -33.949 | -2.3114
1.3613 | -33.9506 | -2.3192
1.3674 | -33.9519 | -2.3277
1.3735 | -33.9532 | -2.3357
1.3796 | -33.9544 | -2.343
1.3857 | -33.9554 | -2.35
1.3918 | -33.9565 | -2.3596
1.3979 | -33.9578 | -2.3681
1.404 | -33.9585 | -2.378
1.4101 | -33.9594 | -2.3869
1.4162 | -33.9602 | -2.3953
1.4223 | -33.961 | -2.4056
1.4284 | -33.9618 | -2.4138
1.4345 | -33.9624 | -2.424
1.4406 | -33.9629 | -2.4349
1.4467 | -33.9635 | -2.4442
1.4528 | -33.9642 | -2.4533
1.4589 | -33.965 | -2.4637
1.465 | -33.9652 | -2.4713
1.4711 | -33.9661 | -2.4801
1.4772 | -33.9666 | -2.4921
1.4833 | -33.9672 | -2.5042
1.4894 | -33.9679 | -2.5155
1.4955 | -33.9686 | -2.5268
1.5016 | -33.9692 | -2.5388
1.5078 | -33.9696 | -2.5531
1.5139 | -33.9703 | -2.5657
1.52 | -33.9708 | -2.5788
1.5261 | -33.9711 | -2.5903
1.5322 | -33.9713 | -2.6028
1.5383 | -33.9716 | -2.6147
1.5444 | -33.9722 | -2.6279
1.5505 | -33.9726 | -2.641
1.5566 | -33.9731 | -2.6537
1.5627 | -33.9736 | -2.6669
1.5688 | -33.9744 | -2.6798
1.5749 | -33.9746 | -2.692
1.581 | -33.9749 | -2.7085
1.5871 | -33.9754 | -2.724
1.5932 | -33.9758 | -2.7367
1.5993 | -33.9762 | -2.7534
1.6054 | -33.9765 | -2.7681
1.6115 | -33.9768 | -2.7849
1.6176 | -33.9771 | -2.8012
1.6237 | -33.9773 | -2.8154
1.6298 | -33.9776 | -2.8307
1.6359 | -33.978 | -2.845
1.642 | -33.9784 | -2.8594
1.6482 | -33.9788 | -2.8763
1.6543 | -33.9792 | -2.8929
1.6604 | -33.9797 | -2.9074
1.6665 | -33.9797 | -2.9222
1.6726 | -33.98 | -2.9383
1.6787 | -33.9805 | -2.9552
1.6848 | -33.9803 | -2.9715
1.6909 | -33.9807 | -2.9876
1.697 | -33.9809 | -3.0036
1.7031 | -33.9809 | -3.0192
1.7092 | -33.9809 | -3.034
1.7153 | -33.9811 | -3.0501
1.7214 | -33.9807 | -3.0653
1.7275 | -33.9805 | -3.081
1.7336 | -33.9803 | -3.0971
1.7397 | -33.9803 | -3.1122
1.7458 | -33.9804 | -3.1288
1.7519 | -33.9801 | -3.1448
1.758 | -33.9801 | -3.1596
1.7641 | -33.9802 | -3.174
1.7702 | -33.9799 | -3.1897
1.7763 | -33.9799 | -3.2033
1.7824 | -33.9797 | -3.2194
1.7885 | -33.9795 | -3.2333
1.7947 | -33.9796 | -3.248
1.8008 | -33.9795 | -3.2608
1.8069 | -33.9794 | -3.2743
1.813 | -33.9795 | -3.2884
1.8191 | -33.9792 | -3.3026
1.8252 | -33.9789 | -3.3164
1.8313 | -33.9789 | -3.3307
1.8374 | -33.9789 | -3.3441
1.8435 | -33.979 | -3.3562
1.8496 | -33.9787 | -3.3706
1.8557 | -33.9786 | -3.3852
1.8618 | -33.9785 | -3.397
1.8679 | -33.9785 | -3.41
1.874 | -33.9785 | -3.423
1.8801 | -33.9784 | -3.4359
1.8862 | -33.9785 | -3.4487
1.8923 | -33.9785 | -3.4607
1.8984 | -33.9781 | -3.4714
1.9045 | -33.978 | -3.4824
1.9106 | -33.9777 | -3.4941
1.9167 | -33.9774 | -3.5053
1.9228 | -33.9774 | -3.5161
1.9289 | -33.977 | -3.5249
1.9351 | -33.9771 | -3.5339
1.9412 | -33.9772 | -3.5441
1.9473 | -33.9772 | -3.5561
1.9534 | -33.9771 | -3.5671
1.9595 | -33.9769 | -3.5786
1.9656 | -33.9764 | -3.5892
1.9717 | -33.9761 | -3.5997
1.9778 | -33.9761 | -3.6112
1.9839 | -33.9761 | -3.6226
1.99 | -33.9762 | -3.6324
1.9961 | -33.9759 | -3.6427
2.0022 | -33.9758 | -3.6526
2.0083 | -33.9754 | -3.6626
2.0144 | -33.9754 | -3.6724
2.0205 | -33.9754 | -3.6842
2.0266 | -33.9755 | -3.6957
2.0327 | -33.9755 | -3.7068
2.0388 | -33.9757 | -3.7206
2.0449 | -33.9758 | -3.7309
2.051 | -33.9758 | -3.7408
2.0571 | -33.9756 | -3.7521
2.0632 | -33.9755 | -3.7626
2.0693 | -33.9755 | -3.7737
2.0754 | -33.9757 | -3.7854
2.0816 | -33.9758 | -3.7973
2.0877 | -33.9757 | -3.8083
2.0938 | -33.9759 | -3.8206
2.0999 | -33.9758 | -3.8321
2.106 | -33.976 | -3.8439
2.1121 | -33.976 | -3.8561
2.1182 | -33.9761 | -3.8661
2.1243 | -33.9763 | -3.8766
2.1304 | -33.9763 | -3.8867
2.1365 | -33.9767 | -3.8982
2.1426 | -33.9768 | -3.9114
2.1487 | -33.9772 | -3.923
2.1548 | -33.9776 | -3.936
2.1609 | -33.9781 | -3.9479
2.167 | -33.9786 | -3.9604
2.1731 | -33.9789 | -3.9722
2.1792 | -33.9795 | -3.9833
2.1853 | -33.98 | -3.996
2.1914 | -33.9805 | -4.0081
2.1975 | -33.981 | -4.0202
2.2036 | -33.9814 | -4.0333
2.2097 | -33.9818 | -4.0452
2.2158 | -33.9828 | -4.0554
2.222 | -33.9833 | -4.0654
2.2281 | -33.9836 | -4.0761
2.2342 | -33.9841 | -4.0891
2.2403 | -33.9845 | -4.0991
2.2464 | -33.9853 | -4.1101
2.2525 | -33.9857 | -4.1208
2.2586 | -33.9866 | -4.1305
2.2647 | -33.9874 | -4.1398
2.2708 | -33.988 | -4.1514
2.2769 | -33.9888 | -4.1628
2.283 | -33.9893 | -4.1721
2.2891 | -33.9898 | -4.1824
2.2952 | -33.9904 | -4.1921
2.3013 | -33.991 | -4.2022
2.3074 | -33.9917 | -4.2118
2.3135 | -33.9921 | -4.223
2.3196 | -33.9927 | -4.2315
2.3257 | -33.9931 | -4.242
2.3318 | -33.9934 | -4.2509
2.3379 | -33.9939 | -4.2623
2.344 | -33.9942 | -4.2731
2.3501 | -33.9947 | -4.2838
2.3562 | -33.995 | -4.292
2.3623 | -33.9953 | -4.3006
2.3685 | -33.9955 | -4.3096
2.3746 | -33.9957 | -4.32
2.3807 | -33.9962 | -4.3289
2.3868 | -33.9964 | -4.3378
2.3929 | -33.9968 | -4.3471
2.399 | -33.997 | -4.355
2.4051 | -33.9975 | -4.3644
2.4112 | -33.9978 | -4.3734
2.4173 | -33.9981 | -4.3822
2.4234 | -33.9981 | -4.3907
2.4295 | -33.9984 | -4.3992
2.4356 | -33.9985 | -4.4084
2.4417 | -33.9985 | -4.4184
2.4478 | -33.9986 | -4.4279
2.4539 | -33.9989 | -4.4358
2.46 | -33.9985 | -4.4454
2.4661 | -33.9986 | -4.4546
2.4722 | -33.9988 | -4.4633
2.4783 | -33.9989 | -4.4716
2.4844 | -33.9989 | -4.4801
2.4905 | -33.9988 | -4.4892
2.4966 | -33.9985 | -4.498
2.5027 | -33.9985 | -4.506
2.5089 | -33.9984 | -4.5145
2.515 | -33.9981 | -4.5236
2.5211 | -33.9981 | -4.5319
2.5272 | -33.9978 | -4.5409
2.5333 | -33.9976 | -4.5504
2.5394 | -33.9975 | -4.5593
2.5455 | -33.9973 | -4.5686
2.5516 | -33.9973 | -4.5773
2.5577 | -33.9972 | -4.5873
2.5638 | -33.9969 | -4.5975
2.5699 | -33.9967 | -4.6077
2.576 | -33.9965 | -4.6181
2.5821 | -33.9963 | -4.628
2.5882 | -33.9961 | -4.6375
2.5943 | -33.9959 | -4.6472
2.6004 | -33.9957 | -4.6571
2.6065 | -33.9955 | -4.6655
2.6126 | -33.9954 | -4.6746
2.6187 | -33.9953 | -4.6832
2.6248 | -33.9952 | -4.6922
2.6309 | -33.995 | -4.7034
2.637 | -33.995 | -4.7145
2.6431 | -33.995 | -4.724
2.6492 | -33.9948 | -4.7339
2.6554 | -33.9948 | -4.7441
2.6615 | -33.9946 | -4.754
2.6676 | -33.9946 | -4.7643
2.6737 | -33.9947 | -4.775
2.6798 | -33.9947 | -4.7849
2.6859 | -33.9947 | -4.7953
2.692 | -33.9946 | -4.805
2.6981 | -33.9945 | -4.8161
2.7042 | -33.9946 | -4.8281
2.7103 | -33.9945 | -4.8381
2.7164 | -33.9948 | -4.8488
2.7225 | -33.9949 | -4.8601
2.7286 | -33.9949 | -4.8712
2.7347 | -33.995 | -4.8811
2.7408 | -33.9952 | -4.8903
2.7469 | -33.9952 | -4.9016
2.753 | -33.9955 | -4.9122
2.7591 | -33.9956 | -4.9241
2.7652 | -33.9956 | -4.9356
2.7713 | -33.9956 | -4.9454
2.7774 | -33.996 | -4.9568
2.7835 | -33.9961 | -4.9674
2.7896 | -33.9963 | -4.9766
2.7958 | -33.9967 | -4.9869
2.8019 | -33.9972 | -4.9966
2.808 | -33.9975 | -5.0084
2.8141 | -33.9979 | -5.0213
2.8202 | -33.9982 | -5.0334
2.8263 | -33.9987 | -5.0443
2.8324 | -33.9989 | -5.0567
2.8385 | -33.9993 | -5.0687
2.8446 | -33.9999 | -5.0804
2.8507 | -34.0003 | -5.0924
2.8568 | -34.0007 | -5.1036
2.8629 | -34.0011 | -5.1148
2.869 | -34.0016 | -5.1261
2.8751 | -34.002 | -5.138
2.8812 | -34.0026 | -5.1513
2.8873 | -34.003 | -5.1625
2.8934 | -34.0033 | -5.1744
2.8995 | -34.0037 | -5.1863
2.9056 | -34.004 | -5.197
2.9117 | -34.0045 | -5.2085
2.9178 | -34.0052 | -5.2212
2.9239 | -34.0056 | -5.2345
2.93 | -34.006 | -5.2479
2.9361 | -34.0066 | -5.2592
2.9423 | -34.0072 | -5.2712
2.9484 | -34.0077 | -5.2847
2.9545 | -34.0084 | -5.2979
2.9606 | -34.0089 | -5.3112
2.9667 | -34.0094 | -5.323
2.9728 | -34.0098 | -5.3352
2.9789 | -34.0103 | -5.3479
2.985 | -34.0107 | -5.3616
2.9911 | -34.0112 | -5.3741
2.9972 | -34.0117 | -5.3868
3.0033 | -34.0123 | -5.3998
3.0094 | -34.0128 | -5.4128
3.0155 | -34.0134 | -5.4254
3.0216 | -34.0139 | -5.4388
3.0277 | -34.0144 | -5.4525
3.0338 | -34.015 | -5.466
3.0399 | -34.0156 | -5.4797
3.046 | -34.0159 | -5.4927
3.0521 | -34.0164 | -5.5051
3.0582 | -34.0167 | -5.5189
3.0643 | -34.0173 | -5.5319
3.0704 | -34.0175 | -5.5453
3.0765 | -34.0179 | -5.5575
3.0827 | -34.0182 | -5.5705
3.0888 | -34.0186 | -5.5833
3.0949 | -34.0189 | -5.5963
3.101 | -34.0193 | -5.6104
3.1071 | -34.0196 | -5.6245
3.1132 | -34.0201 | -5.6385
3.1193 | -34.0205 | -5.6515
3.1254 | -34.0209 | -5.6636
3.1315 | -34.0213 | -5.678
3.1376 | -34.0216 | -5.6908
3.1437 | -34.0219 | -5.7043
3.1498 | -34.0223 | -5.7184
3.1559 | -34.0226 | -5.7316
3.162 | -34.0226 | -5.7457
3.1681 | -34.0227 | -5.7599
3.1742 | -34.023 | -5.7736
3.1803 | -34.0234 | -5.7867
3.1864 | -34.0238 | -5.799
3.1925 | -34.0241 | -5.8126
3.1986 | -34.0241 | -5.8254
3.2047 | -34.0243 | -5.8393
3.2108 | -34.0245 | -5.851
3.2169 | -34.0249 | -5.8647
3.223 | -34.0248 | -5.8781
3.2292 | -34.0248 | -5.892
3.2353 | -34.0247 | -5.9015
3.2414 | -34.0249 | -5.9144
3.2475 | -34.0248 | -5.9254
3.2536 | -34.0251 | -5.9391
3.2597 | -34.0253 | -5.9503
3.2658 | -34.0255 | -5.961
3.2719 | -34.0256 | -5.972
3.278 | -34.0256 | -5.9841
3.2841 | -34.0258 | -5.996
3.2902 | -34.0259 | -6.0085
3.2963 | -34.026 | -6.0202
3.3024 | -34.0258 | -6.0303
3.3085 | -34.0261 | -6.0428
3.3146 | -34.026 | -6.0537
3.3207 | -34.0262 | -6.0658
3.3268 | -34.0263 | -6.078
3.3329 | -34.0264 | -6.0877
3.339 | -34.0264 | -6.0974
3.3451 | -34.0265 | -6.1087
3.3512 | -34.0266 | -6.1181
3.3573 | -34.0267 | -6.1294
3.3634 | -34.0268 | -6.1409
3.3696 | -34.0267 | -6.1528
3.3757 | -34.0267 | -6.1626
3.3818 | -34.027 | -6.1743
3.3879 | -34.0269 | -6.1846
3.394 | -34.0269 | -6.1951
3.4001 | -34.0269 | -6.2059
3.4062 | -34.0269 | -6.2161
3.4123 | -34.0269 | -6.2258
3.4184 | -34.0269 | -6.2368
3.4245 | -34.027 | -6.2473
3.4306 | -34.027 | -6.257
3.4367 | -34.0271 | -6.2667
3.4428 | -34.0274 | -6.2758
3.4489 | -34.0273 | -6.2869
3.455 | -34.0272 | -6.2973
3.4611 | -34.0272 | -6.3084
3.4672 | -34.0272 | -6.3196
3.4733 | -34.0272 | -6.3298
3.4794 | -34.0272 | -6.3403
3.4855 | -34.0274 | -6.3498
3.4916 | -34.0275 | -6.3595
3.4977 | -34.0275 | -6.3704
3.5038 | -34.0276 | -6.3803
3.5099 | -34.0277 | -6.3899
3.5161 | -34.0278 | -6.3994
3.5222 | -34.0279 | -6.4094
3.5283 | -34.0278 | -6.4194
3.5344 | -34.0278 | -6.4302
3.5405 | -34.0279 | -6.4401
3.5466 | -34.028 | -6.4498
3.5527 | -34.0281 | -6.4605
3.5588 | -34.0282 | -6.4702
3.5649 | -34.0282 | -6.4803
3.571 | -34.0283 | -6.4905
3.5771 | -34.0284 | -6.5012
3.5832 | -34.0285 | -6.5108
3.5893 | -34.0287 | -6.5206
3.5954 | -34.0288 | -6.5311
3.6015 | -34.0289 | -6.5415
3.6076 | -34.0291 | -6.551
3.6137 | -34.0292 | -6.5609
3.6198 | -34.0295 | -6.5699
3.6259 | -34.0296 | -6.5806
3.632 | -34.0298 | -6.5912
3.6381 | -34.03 | -6.6011
3.6442 | -34.0302 | -6.6115
3.6503 | -34.0304 | -6.6205
3.6565 | -34.0307 | -6.6322
3.6626 | -34.0307 | -6.6429
3.6687 | -34.031 | -6.6535
3.6748 | -34.0313 | -6.6634
3.6809 | -34.0317 | -6.6738
3.687 | -34.0318 | -6.685
3.6931 | -34.0322 | -6.695
3.6992 | -34.0324 | -6.7056
3.7053 | -34.0326 | -6.7157
3.7114 | -34.0328 | -6.7269
3.7175 | -34.0332 | -6.7368
3.7236 | -34.0333 | -6.7455
3.7297 | -34.0334 | -6.7562
3.7358 | -34.0337 | -6.7679
3.7419 | -34.0339 | -6.7795
3.748 | -34.034 | -6.7902
3.7541 | -34.0341 | -6.8012
3.7602 | -34.0345 | -6.812
3.7663 | -34.0347 | -6.8221
3.7724 | -34.035 | -6.8331
3.7785 | -34.0353 | -6.8435
3.7846 | -34.0355 | -6.8541
3.7907 | -34.0359 | -6.8651
3.7969 | -34.0361 | -6.8766
3.803 | -34.0364 | -6.8866
3.8091 | -34.0367 | -6.8969
3.8152 | -34.037 | -6.9081
3.8213 | -34.0372 | -6.9197
3.8274 | -34.0376 | -6.9301
3.8335 | -34.038 | -6.9417
3.8396 | -34.0384 | -6.9532
3.8457 | -34.0387 | -6.9634
3.8518 | -34.0392 | -6.9742
3.8579 | -34.0396 | -6.9848
3.864 | -34.0399 | -6.9955
3.8701 | -34.0403 | -7.0064
3.8762 | -34.0406 | -7.0176
3.8823 | -34.0409 | -7.0282
3.8884 | -34.0413 | -7.0385
3.8945 | -34.0418 | -7.0499
3.9006 | -34.0422 | -7.0614
3.9067 | -34.0425 | -7.0732
3.9128 | -34.0429 | -7.0845
3.9189 | -34.0433 | -7.0951
3.925 | -34.0435 | -7.1059
3.9311 | -34.0439 | -7.1159
3.9372 | -34.0443 | -7.1274
3.9434 | -34.0446 | -7.1378
3.9495 | -34.045 | -7.1494
3.9556 | -34.0454 | -7.1606
3.9617 | -34.0458 | -7.1715
3.9678 | -34.0462 | -7.1821
3.9739 | -34.0466 | -7.1933
3.98 | -34.0469 | -7.2051
3.9861 | -34.0472 | -7.2155
3.9922 | -34.0474 | -7.2272
3.9983 | -34.0478 | -7.2372
4.0044 | -34.0481 | -7.2469
4.0105 | -34.0485 | -7.2578
4.0166 | -34.049 | -7.2688
4.0227 | -34.0494 | -7.2791
4.0288 | -34.0496 | -7.2905
4.0349 | -34.0499 | -7.3003
4.041 | -34.0503 | -7.3112
4.0471 | -34.0505 | -7.3214
4.0532 | -34.051 | -7.3325
4.0593 | -34.0512 | -7.3441
4.0654 | -34.0516 | -7.3548
4.0715 | -34.0519 | -7.3661
4.0776 | -34.0522 | -7.3765
4.0838 | -34.0526 | -7.3878
4.0899 | -34.0529 | -7.3991
4.096 | -34.0532 | -7.4095
4.1021 | -34.0535 | -7.4202
4.1082 | -34.0539 | -7.4304
4.1143 | -34.0542 | -7.4404
4.1204 | -34.0547 | -7.4511
4.1265 | -34.055 | -7.4614
4.1326 | -34.0554 | -7.4717
4.1387 | -34.0558 | -7.4829
4.1448 | -34.0562 | -7.4938
4.1509 | -34.0564 | -7.5045
4.157 | -34.0567 | -7.5156
4.1631 | -34.057 | -7.5268
4.1692 | -34.0573 | -7.5375
4.1753 | -34.0577 | -7.5483
4.1814 | -34.0579 | -7.558
4.1875 | -34.0582 | -7.5694
4.1936 | -34.0585 | -7.5807
4.1997 | -34.0588 | -7.5908
4.2058 | -34.0591 | -7.6023
4.2119 | -34.0594 | -7.6119
4.218 | -34.0596 | -7.622
4.2241 | -34.06 | -7.6316
4.2303 | -34.0604 | -7.6428
4.2364 | -34.0607 | -7.6542
4.2425 | -34.061 | -7.6647
4.2486 | -34.0612 | -7.6758
4.2547 | -34.0614 | -7.685
4.2608 | -34.0616 | -7.6948
4.2669 | -34.0619 | -7.7052
4.273 | -34.0622 | -7.7159
4.2791 | -34.0624 | -7.7266
4.2852 | -34.0626 | -7.7368
4.2913 | -34.0628 | -7.746
4.2974 | -34.063 | -7.7569
4.3035 | -34.0632 | -7.7676
4.3096 | -34.0634 | -7.778
4.3157 | -34.0636 | -7.7879
4.3218 | -34.0638 | -7.7978
4.3279 | -34.0641 | -7.8082
4.334 | -34.0643 | -7.818
4.3401 | -34.0645 | -7.8278
4.3462 | -34.0648 | -7.8386
4.3523 | -34.0651 | -7.8483
4.3584 | -34.0652 | -7.8567
4.3645 | -34.0654 | -7.8663
4.3707 | -34.0656 | -7.8755
4.3768 | -34.0658 | -7.8859
4.3829 | -34.066 | -7.8961
4.389 | -34.0662 | -7.9058
4.3951 | -34.0665 | -7.9153
4.4012 | -34.0667 | -7.925
4.4073 | -34.0669 | -7.9351
4.4134 | -34.0671 | -7.9452
4.4195 | -34.0673 | -7.9548
4.4256 | -34.0676 | -7.9643
4.4317 | -34.0678 | -7.9738
4.4378 | -34.068 | -7.9833
4.4439 | -34.0682 | -7.9927
4.45 | -34.0683 | -8.0026
4.4561 | -34.0684 | -8.0124
4.4622 | -34.0685 | -8.022
4.4683 | -34.0687 | -8.0311
4.4744 | -34.0689 | -8.0413
4.4805 | -34.0691 | -8.0503
4.4866 | -34.0694 | -8.0595
4.4927 | -34.0695 | -8.0697
4.4988 | -34.0697 | -8.079
4.5049 | -34.0699 | -8.0881
4.511 | -34.0702 | -8.0971
4.5172 | -34.0705 | -8.1069
4.5233 | -34.0706 | -8.1174
4.5294 | -34.0709 | -8.1267
4.5355 | -34.0711 | -8.136
4.5416 | -34.0713 | -8.1451
4.5477 | -34.0715 | -8.1543
4.5538 | -34.0717 | -8.163
4.5599 | -34.072 | -8.1725
4.566 | -34.0722 | -8.1821
4.5721 | -34.0724 | -8.1913
4.5782 | -34.0727 | -8.2011
4.5843 | -34.0729 | -8.2106
4.5904 | -34.0732 | -8.2199
4.5965 | -34.0735 | -8.2285
4.6026 | -34.0737 | -8.2391
4.6087 | -34.0739 | -8.2481
4.6148 | -34.0741 | -8.2587
4.6209 | -34.0745 | -8.2671
4.627 | -34.0748 | -8.2771
4.6331 | -34.0749 | -8.2859
4.6392 | -34.0751 | -8.2951
4.6453 | -34.0754 | -8.3044
4.6514 | -34.0756 | -8.3139
4.6576 | -34.0757 | -8.3244
4.6637 | -34.0759 | -8.3319
4.6698 | -34.0761 | -8.3405
4.6759 | -34.0762 | -8.3502
4.682 | -34.0764 | -8.3593
4.6881 | -34.0765 | -8.368
4.6942 | -34.0768 | -8.3772
4.7003 | -34.077 | -8.3866
4.7064 | -34.0772 | -8.3959
4.7125 | -34.0776 | -8.4053
4.7186 | -34.0777 | -8.4136
4.7247 | -34.078 | -8.4232
4.7308 | -34.0783 | -8.4325
4.7369 | -34.0785 | -8.442
4.743 | -34.0787 | -8.4523
4.7491 | -34.0789 | -8.4619
4.7552 | -34.079 | -8.4711
4.7613 | -34.0792 | -8.4806
4.7674 | -34.0794 | -8.4892
4.7735 | -34.0795 | -8.4992
4.7796 | -34.0798 | -8.509
4.7857 | -34.0799 | -8.5186
4.7918 | -34.0801 | -8.5272
4.7979 | -34.0804 | -8.5362
4.8041 | -34.0805 | -8.5454
4.8102 | -34.0808 | -8.5544
4.8163 | -34.0811 | -8.5639
4.8224 | -34.0811 | -8.5725
4.8285 | -34.0815 | -8.5815
4.8346 | -34.0817 | -8.591
4.8407 | -34.0819 | -8.6003
4.8468 | -34.0821 | -8.6099
4.8529 | -34.0823 | -8.6197
4.859 | -34.0826 | -8.6292
4.8651 | -34.0828 | -8.6385
4.8712 | -34.083 | -8.6478
4.8773 | -34.0832 | -8.6574
4.8834 | -34.0835 | -8.6671
4.8895 | -34.0837 | -8.6771
4.8956 | -34.0839 | -8.687
4.9017 | -34.084 | -8.6965
4.9078 | -34.0842 | -8.7063
4.9139 | -34.0845 | -8.7148
4.92 | -34.0847 | -8.7238
4.9261 | -34.085 | -8.7326
4.9322 | -34.0854 | -8.7417
4.9383 | -34.0857 | -8.7507
4.9445 | -34.086 | -8.7608
4.9506 | -34.0862 | -8.7696
4.9567 | -34.0864 | -8.7788
4.9628 | -34.0867 | -8.7882
4.9689 | -34.0868 | -8.7979
4.975 | -34.087 | -8.8072
4.9811 | -34.0873 | -8.817
4.9872 | -34.0876 | -8.8268
4.9933 | -34.0879 | -8.8365
4.9994 | -34.0883 | -8.8463
5.0055 | -34.0886 | -8.8559
5.0116 | -34.0889 | -8.8662
5.0177 | -34.0893 | -8.8762
5.0238 | -34.0896 | -8.8867
5.0299 | -34.09 | -8.8977
5.036 | -34.0903 | -8.9072
5.0421 | -34.0905 | -8.9168
5.0482 | -34.0907 | -8.9271
5.0543 | -34.0911 | -8.9364
5.0604 | -34.0914 | -8.9466
5.0665 | -34.0918 | -8.9563
5.0726 | -34.092 | -8.9667
5.0787 | -34.0924 | -8.9769
5.0848 | -34.0927 | -8.9869
5.091 | -34.0929 | -8.9968
5.0971 | -34.0933 | -9.0069
5.1032 | -34.0936 | -9.0172
5.1093 | -34.0939 | -9.0271
5.1154 | -34.0942 | -9.0371
5.1215 | -34.0944 | -9.0471
5.1276 | -34.0947 | -9.0569
5.1337 | -34.0952 | -9.0666
5.1398 | -34.0954 | -9.0764
5.1459 | -34.0958 | -9.0868
5.152 | -34.0961 | -9.0979
5.1581 | -34.0963 | -9.1085
5.1642 | -34.0966 | -9.1185
5.1703 | -34.0969 | -9.1285
5.1764 | -34.0972 | -9.1384
5.1825 | -34.0975 | -9.1487
5.1886 | -34.0978 | -9.1585
5.1947 | -34.0981 | -9.1693
5.2008 | -34.0985 | -9.1791
5.2069 | -34.0988 | -9.1903
5.213 | -34.0991 | -9.1998
5.2191 | -34.0995 | -9.2102
5.2252 | -34.0998 | -9.2205
5.2314 | -34.1 | -9.2312
5.2375 | -34.1004 | -9.2411
5.2436 | -34.1007 | -9.2513
5.2497 | -34.101 | -9.2612
5.2558 | -34.1014 | -9.2722
5.2619 | -34.1017 | -9.283
5.268 | -34.102 | -9.2935
5.2741 | -34.1024 | -9.3047
5.2802 | -34.1027 | -9.3155
5.2863 | -34.1031 | -9.3257
5.2924 | -34.1034 | -9.3369
5.2985 | -34.1038 | -9.3483
5.3046 | -34.1042 | -9.3595
5.3107 | -34.1046 | -9.3704
5.3168 | -34.105 | -9.3808
5.3229 | -34.1054 | -9.3906
5.329 | -34.1057 | -9.4006
5.3351 | -34.1061 | -9.4108
5.3412 | -34.1065 | -9.4209
5.3473 | -34.1069 | -9.4318
5.3534 | -34.1073 | -9.4427
5.3595 | -34.1076 | -9.4534
5.3656 | -34.108 | -9.4644
5.3717 | -34.1083 | -9.475
5.3779 | -34.1086 | -9.4859
5.384 | -34.1089 | -9.4966
5.3901 | -34.1093 | -9.5073
5.3962 | -34.1096 | -9.5181
5.4023 | -34.11 | -9.5289
5.4084 | -34.1104 | -9.5393
5.4145 | -34.1107 | -9.5504
5.4206 | -34.1109 | -9.5617
5.4267 | -34.1113 | -9.573
5.4328 | -34.1117 | -9.5832
5.4389 | -34.112 | -9.5934
5.445 | -34.1123 | -9.6048
5.4511 | -34.1127 | -9.6152
5.4572 | -34.1131 | -9.6264
5.4633 | -34.1134 | -9.6367
5.4694 | -34.1137 | -9.6472
5.4755 | -34.1141 | -9.6581
5.4816 | -34.1144 | -9.6687
5.4877 | -34.1147 | -9.6794
5.4938 | -34.1151 | -9.6899
5.4999 | -34.1154 | -9.7005
5.506 | -34.1158 | -9.7113
5.5121 | -34.1161 | -9.7224
5.5183 | -34.1164 | -9.7325
5.5244 | -34.1166 | -9.7433
5.5305 | -34.117 | -9.7541
5.5366 | -34.1174 | -9.7646
5.5427 | -34.1177 | -9.7758
5.5488 | -34.118 | -9.7869
5.5549 | -34.1183 | -9.7976
5.561 | -34.1186 | -9.8089
5.5671 | -34.119 | -9.8193
5.5732 | -34.1193 | -9.8302
5.5793 | -34.1196 | -9.8418
5.5854 | -34.1201 | -9.8527
5.5915 | -34.1203 | -9.8636
5.5976 | -34.1206 | -9.8747
5.6037 | -34.1209 | -9.8854
5.6098 | -34.1212 | -9.8959
5.6159 | -34.1216 | -9.9068
5.622 | -34.1219 | -9.9173
5.6281 | -34.1223 | -9.9281
5.6342 | -34.1226 | -9.939
5.6403 | -34.1229 | -9.9493
5.6464 | -34.1233 | -9.9601
5.6525 | -34.1235 | -9.9705
5.6586 | -34.1239 | -9.9805
5.6648 | -34.1242 | -9.9918
5.6709 | -34.1246 | -10.003
5.677 | -34.1249 | -10.0144
5.6831 | -34.1252 | -10.0247
5.6892 | -34.1255 | -10.0355
5.6953 | -34.1258 | -10.0461
5.7014 | -34.1262 | -10.0574
5.7075 | -34.1264 | -10.0684
5.7136 | -34.1268 | -10.0789
5.7197 | -34.1271 | -10.0895
5.7258 | -34.1273 | -10.1001
5.7319 | -34.1277 | -10.1108
5.738 | -34.128 | -10.1215
5.7441 | -34.1283 | -10.1322
5.7502 | -34.1287 | -10.1425
5.7563 | -34.1291 | -10.1529
5.7624 | -34.1294 | -10.1644
5.7685 | -34.1296 | -10.1749
5.7746 | -34.1298 | -10.1854
5.7807 | -34.1301 | -10.1957
5.7868 | -34.1304 | -10.2064
5.7929 | -34.1307 | -10.217
5.799 | -34.131 | -10.2275
5.8052 | -34.1313 | -10.2383
5.8113 | -34.1316 | -10.249
5.8174 | -34.1318 | -10.2598
5.8235 | -34.1321 | -10.2702
5.8296 | -34.1324 | -10.2815
5.8357 | -34.1326 | -10.2922
5.8418 | -34.133 | -10.303
5.8479 | -34.1332 | -10.3139
5.854 | -34.1335 | -10.3249
5.8601 | -34.1338 | -10.3353
5.8662 | -34.1341 | -10.3463
5.8723 | -34.1343 | -10.3563
5.8784 | -34.1346 | -10.3669
5.8845 | -34.135 | -10.3777
5.8906 | -34.1352 | -10.3883
5.8967 | -34.1355 | -10.3985
5.9028 | -34.1359 | -10.4092
5.9089 | -34.1361 | -10.4192
5.915 | -34.1364 | -10.4298
5.9211 | -34.1367 | -10.4404
5.9272 | -34.137 | -10.4509
5.9333 | -34.1373 | -10.4608
5.9394 | -34.1377 | -10.4712
5.9455 | -34.1379 | -10.4813
5.9517 | -34.1383 | -10.4919
5.9578 | -34.1386 | -10.5028
5.9639 | -34.139 | -10.5129
5.97 | -34.1393 | -10.5236
5.9761 | -34.1396 | -10.5338
5.9822 | -34.1399 | -10.5445
5.9883 | -34.1403 | -10.5548
5.9944 | -34.1406 | -10.5653
6.0005 | -34.141 | -10.5761
6.0066 | -34.1412 | -10.5862
6.0127 | -34.1415 | -10.596
6.0188 | -34.1418 | -10.6069
6.0249 | -34.1421 | -10.6173
6.031 | -34.1424 | -10.6275
6.0371 | -34.1428 | -10.6382
6.0432 | -34.1432 | -10.6487
6.0493 | -34.1435 | -10.6589
6.0554 | -34.1439 | -10.6693
6.0615 | -34.1442 | -10.6803
6.0676 | -34.1445 | -10.6913
6.0737 | -34.1448 | -10.702
6.0798 | -34.1451 | -10.7123
6.0859 | -34.1454 | -10.7227
6.0921 | -34.1458 | -10.7329
6.0982 | -34.1461 | -10.7436
6.1043 | -34.1465 | -10.7539
6.1104 | -34.1469 | -10.7648
6.1165 | -34.1472 | -10.7745
6.1226 | -34.1475 | -10.7843
6.1287 | -34.1479 | -10.7948
6.1348 | -34.1483 | -10.8053
6.1409 | -34.1486 | -10.8156
6.147 | -34.149 | -10.8258
6.1531 | -34.1493 | -10.8362
6.1592 | -34.1496 | -10.8466
6.1653 | -34.1499 | -10.8567
6.1714 | -34.1503 | -10.8672
6.1775 | -34.1506 | -10.8779
6.1836 | -34.1509 | -10.888
6.1897 | -34.1512 | -10.8984
6.1958 | -34.1515 | -10.909
6.2019 | -34.1519 | -10.9196
6.208 | -34.1522 | -10.9304
6.2141 | -34.1526 | -10.941
6.2202 | -34.153 | -10.9518
6.2263 | -34.1533 | -10.9621
6.2325 | -34.1537 | -10.972
6.2386 | -34.1541 | -10.982
6.2447 | -34.1544 | -10.9926
6.2508 | -34.1547 | -11.0029
6.2569 | -34.1551 | -11.0136
6.263 | -34.1554 | -11.0236
6.2691 | -34.1558 | -11.0339
6.2752 | -34.1562 | -11.0441
6.2813 | -34.1565 | -11.0544
6.2874 | -34.1568 | -11.0648
6.2935 | -34.1572 | -11.0749
6.2996 | -34.1575 | -11.0853
6.3057 | -34.1579 | -11.0956
6.3118 | -34.1582 | -11.1059
6.3179 | -34.1586 | -11.1162
6.324 | -34.159 | -11.1267
6.3301 | -34.1593 | -11.1369
6.3362 | -34.1597 | -11.1473
6.3423 | -34.1601 | -11.1577
6.3484 | -34.1605 | -11.1682
6.3545 | -34.1609 | -11.1779
6.3606 | -34.1613 | -11.1879
6.3667 | -34.1617 | -11.198
6.3728 | -34.162 | -11.2083
6.379 | -34.1624 | -11.2186
6.3851 | -34.1627 | -11.229
6.3912 | -34.1631 | -11.239
6.3973 | -34.1635 | -11.249
6.4034 | -34.1639 | -11.2592
6.4095 | -34.1642 | -11.2698
6.4156 | -34.1645 | -11.2798
6.4217 | -34.1649 | -11.29
6.4278 | -34.1652 | -11.3002
6.4339 | -34.1657 | -11.3104
6.44 | -34.1661 | -11.3203
6.4461 | -34.1664 | -11.331
6.4522 | -34.1668 | -11.3407
6.4583 | -34.1671 | -11.351
6.4644 | -34.1673 | -11.3616
6.4705 | -34.1678 | -11.372
6.4766 | -34.1681 | -11.3827
6.4827 | -34.1685 | -11.3934
6.4888 | -34.1688 | -11.4035
6.4949 | -34.1692 | -11.4133
6.501 | -34.1695 | -11.4239
6.5071 | -34.1699 | -11.4343
6.5132 | -34.1704 | -11.4444
6.5194 | -34.1707 | -11.4545
6.5255 | -34.1711 | -11.4644
6.5316 | -34.1714 | -11.4748
6.5377 | -34.1718 | -11.4853
6.5438 | -34.1721 | -11.4955
6.5499 | -34.1726 | -11.5055
6.556 | -34.1729 | -11.5156
6.5621 | -34.1732 | -11.526
6.5682 | -34.1736 | -11.5361
6.5743 | -34.1739 | -11.5471
6.5804 | -34.1742 | -11.5573
6.5865 | -34.1745 | -11.5679
6.5926 | -34.175 | -11.578
6.5987 | -34.1753 | -11.5879
6.6048 | -34.1757 | -11.5988
6.6109 | -34.1761 | -11.609
6.617 | -34.1764 | -11.6201
6.6231 | -34.1766 | -11.6306
6.6292 | -34.1771 | -11.6408
6.6353 | -34.1775 | -11.652
6.6414 | -34.1779 | -11.6626
6.6475 | -34.1783 | -11.6731
6.6536 | -34.1786 | -11.683
6.6597 | -34.1791 | -11.693
6.6659 | -34.1794 | -11.7031
6.672 | -34.1798 | -11.7134
6.6781 | -34.1801 | -11.7235
6.6842 | -34.1805 | -11.7339
6.6903 | -34.1809 | -11.7441
6.6964 | -34.1813 | -11.7545
6.7025 | -34.1816 | -11.7647
6.7086 | -34.182 | -11.7748
6.7147 | -34.1823 | -11.7852
6.7208 | -34.1827 | -11.796
6.7269 | -34.183 | -11.8059
6.733 | -34.1833 | -11.8154
6.7391 | -34.1838 | -11.8259
6.7452 | -34.1842 | -11.8355
6.7513 | -34.1846 | -11.8454
6.7574 | -34.185 | -11.856
6.7635 | -34.1854 | -11.8657
6.7696 | -34.1858 | -11.876
6.7757 | -34.1862 | -11.8866
6.7818 | -34.1866 | -11.8969
6.7879 | -34.1869 | -11.9072
6.794 | -34.1873 | -11.9173
6.8001 | -34.1876 | -11.9273
6.8063 | -34.188 | -11.9377
6.8124 | -34.1884 | -11.948
6.8185 | -34.1887 | -11.9582
6.8246 | -34.1891 | -11.9685
6.8307 | -34.1894 | -11.9786
6.8368 | -34.1898 | -11.9891
6.8429 | -34.1903 | -11.9993
6.849 | -34.1906 | -12.0095
6.8551 | -34.191 | -12.0197
6.8612 | -34.1914 | -12.0298
6.8673 | -34.1917 | -12.0398
6.8734 | -34.1921 | -12.0507
6.8795 | -34.1925 | -12.0612
6.8856 | -34.1929 | -12.0709
6.8917 | -34.1932 | -12.0806
6.8978 | -34.1936 | -12.0907
6.9039 | -34.1939 | -12.1007
6.91 | -34.1943 | -12.111
6.9161 | -34.1947 | -12.121
6.9222 | -34.195 | -12.1312
6.9283 | -34.1954 | -12.1411
6.9344 | -34.1957 | -12.1511
6.9405 | -34.1961 | -12.161
6.9466 | -34.1965 | -12.1715
6.9528 | -34.1968 | -12.1817
6.9589 | -34.1972 | -12.1917
6.965 | -34.1976 | -12.2019
6.9711 | -34.198 | -12.2121
6.9772 | -34.1984 | -12.222
6.9833 | -34.1987 | -12.2323
6.9894 | -34.199 | -12.2429
6.9955 | -34.1994 | -12.253
7.0016 | -34.1997 | -12.2636
7.0077 | -34.2001 | -12.274
7.0138 | -34.2005 | -12.2845
7.0199 | -34.2008 | -12.2945
7.026 | -34.2011 | -12.3045
7.0321 | -34.2015 | -12.3147
7.0382 | -34.2018 | -12.3246
7.0443 | -34.2022 | -12.3351
7.0504 | -34.2025 | -12.3456
7.0565 | -34.2029 | -12.3559
7.0626 | -34.2032 | -12.3659
7.0687 | -34.2036 | -12.3757
7.0748 | -34.204 | -12.3856
7.0809 | -34.2043 | -12.3959
7.087 | -34.2046 | -12.4061
7.0932 | -34.205 | -12.4162
7.0993 | -34.2053 | -12.4264
7.1054 | -34.2056 | -12.4363
7.1115 | -34.206 | -12.4463
7.1176 | -34.2062 | -12.4562
7.1237 | -34.2065 | -12.4663
7.1298 | -34.2068 | -12.476
7.1359 | -34.2072 | -12.4858
7.142 | -34.2075 | -12.496
7.1481 | -34.2078 | -12.5058
7.1542 | -34.2081 | -12.5157
7.1603 | -34.2084 | -12.5257
7.1664 | -34.2088 | -12.5358
7.1725 | -34.2092 | -12.5456
7.1786 | -34.2096 | -12.5558
7.1847 | -34.2099 | -12.5657
7.1908 | -34.2102 | -12.5759
7.1969 | -34.2105 | -12.5859
7.203 | -34.2109 | -12.596
7.2091 | -34.2112 | -12.6058
7.2152 | -34.2115 | -12.6158
7.2213 | -34.2118 | -12.6256
7.2274 | -34.2122 | -12.6356
7.2335 | -34.2125 | -12.6456
7.2397 | -34.2129 | -12.6556
7.2458 | -34.2132 | -12.6652
7.2519 | -34.2135 | -12.6751
7.258 | -34.2138 | -12.6853
7.2641 | -34.2142 | -12.6953
7.2702 | -34.2145 | -12.7056
7.2763 | -34.2149 | -12.7153
7.2824 | -34.2153 | -12.7252
7.2885 | -34.2156 | -12.7351
7.2946 | -34.216 | -12.7453
7.3007 | -34.2162 | -12.7553
7.3068 | -34.2165 | -12.7655
7.3129 | -34.2169 | -12.7753
7.319 | -34.2173 | -12.7854
7.3251 | -34.2176 | -12.795
7.3312 | -34.2179 | -12.8046
7.3373 | -34.2183 | -12.8143
7.3434 | -34.2187 | -12.824
7.3495 | -34.2191 | -12.8335
7.3556 | -34.2194 | -12.8438
7.3617 | -34.2197 | -12.8542
7.3678 | -34.2201 | -12.8642
7.3739 | -34.2204 | -12.8737
7.3801 | -34.2207 | -12.8836
7.3862 | -34.221 | -12.8934
7.3923 | -34.2213 | -12.9033
7.3984 | -34.2217 | -12.9131
7.4045 | -34.222 | -12.9229
7.4106 | -34.2224 | -12.9327
7.4167 | -34.2227 | -12.9428
7.4228 | -34.223 | -12.9526
7.4289 | -34.2234 | -12.9622
7.435 | -34.2238 | -12.9718
7.4411 | -34.2241 | -12.9814
7.4472 | -34.2244 | -12.9914
7.4533 | -34.2248 | -13.0013
7.4594 | -34.2251 | -13.0115
7.4655 | -34.2255 | -13.0213
7.4716 | -34.2258 | -13.0314
7.4777 | -34.2261 | -13.0407
7.4838 | -34.2265 | -13.0507
7.4899 | -34.2269 | -13.0606
7.496 | -34.2272 | -13.0705
7.5021 | -34.2276 | -13.0802
7.5082 | -34.2279 | -13.0901
7.5143 | -34.2282 | -13.0998
7.5204 | -34.2285 | -13.1094
7.5266 | -34.2288 | -13.1193
7.5327 | -34.2292 | -13.129
7.5388 | -34.2295 | -13.1387
7.5449 | -34.2298 | -13.1488
7.551 | -34.2301 | -13.1588
7.5571 | -34.2304 | -13.1685
7.5632 | -34.2308 | -13.1783
7.5693 | -34.2311 | -13.1882
7.5754 | -34.2315 | -13.1981
7.5815 | -34.2319 | -13.2075
7.5876 | -34.2321 | -13.2173
7.5937 | -34.2324 | -13.227
7.5998 | -34.2328 | -13.2369
7.6059 | -34.2331 | -13.2467
7.612 | -34.2334 | -13.2564
7.6181 | -34.2337 | -13.2662
7.6242 | -34.2341 | -13.2761
7.6303 | -34.2345 | -13.2857
7.6364 | -34.2348 | -13.2957
7.6425 | -34.2351 | -13.3058
7.6486 | -34.2354 | -13.3157
7.6547 | -34.2358 | -13.3256
7.6608 | -34.2362 | -13.3353
7.667 | -34.2365 | -13.3446
7.6731 | -34.2368 | -13.3545
7.6792 | -34.2371 | -13.3645
7.6853 | -34.2374 | -13.3742
7.6914 | -34.2376 | -13.3841
7.6975 | -34.238 | -13.3941
7.7036 | -34.2383 | -13.404
7.7097 | -34.2386 | -13.4137
7.7158 | -34.239 | -13.4234
7.7219 | -34.2392 | -13.4331
7.728 | -34.2395 | -13.443
7.7341 | -34.2398 | -13.4529
7.7402 | -34.2402 | -13.4626
7.7463 | -34.2405 | -13.4724
7.7524 | -34.2408 | -13.4825
7.7585 | -34.2411 | -13.4925
7.7646 | -34.2414 | -13.502
7.7707 | -34.2417 | -13.5116
7.7768 | -34.2419 | -13.5213
7.7829 | -34.2422 | -13.5312
7.789 | -34.2425 | -13.5407
7.7951 | -34.2428 | -13.5508
7.8012 | -34.2431 | -13.561
7.8073 | -34.2434 | -13.5707
7.8135 | -34.2437 | -13.5806
7.8196 | -34.244 | -13.5903
7.8257 | -34.2444 | -13.6005
7.8318 | -34.2447 | -13.6107
7.8379 | -34.245 | -13.62
7.844 | -34.2452 | -13.6294
7.8501 | -34.2455 | -13.6395
7.8562 | -34.2457 | -13.6499
7.8623 | -34.2461 | -13.6598
7.8684 | -34.2464 | -13.6693
7.8745 | -34.2467 | -13.6789
7.8806 | -34.2469 | -13.6882
7.8867 | -34.2473 | -13.6979
7.8928 | -34.2476 | -13.7075
7.8989 | -34.248 | -13.7169
7.905 | -34.2483 | -13.727
7.9111 | -34.2486 | -13.7368
7.9172 | -34.2489 | -13.7464
7.9233 | -34.2492 | -13.7563
7.9294 | -34.2495 | -13.7662
7.9355 | -34.2498 | -13.7755
7.9416 | -34.2501 | -13.7852
7.9477 | -34.2504 | -13.7945
7.9539 | -34.2507 | -13.8039
7.96 | -34.251 | -13.814
7.9661 | -34.2513 | -13.8238
7.9722 | -34.2517 | -13.8336
7.9783 | -34.2519 | -13.8433
7.9844 | -34.2523 | -13.8536
7.9905 | -34.2526 | -13.8632
7.9966 | -34.2528 | -13.873
8.0027 | -34.2531 | -13.883
8.0088 | -34.2534 | -13.8924
8.0149 | -34.2537 | -13.9019
8.021 | -34.254 | -13.9117
8.0271 | -34.2543 | -13.9219
8.0332 | -34.2546 | -13.9316
8.0393 | -34.2549 | -13.9415
8.0454 | -34.2552 | -13.9515
8.0515 | -34.2555 | -13.9612
8.0576 | -34.2558 | -13.9712
8.0637 | -34.2561 | -13.9804
8.0698 | -34.2564 | -13.9898
8.0759 | -34.2567 | -13.9992
8.082 | -34.257 | -14.0091
8.0881 | -34.2573 | -14.0187
8.0942 | -34.2576 | -14.0289
8.1004 | -34.2578 | -14.0391
8.1065 | -34.2582 | -14.0488
8.1126 | -34.2585 | -14.0584
8.1187 | -34.2588 | -14.0683
8.1248 | -34.2591 | -14.0782
8.1309 | -34.2594 | -14.088
8.137 | -34.2597 | -14.0976
8.1431 | -34.26 | -14.1075
8.1492 | -34.2603 | -14.1173
8.1553 | -34.2606 | -14.1271
8.1614 | -34.2609 | -14.1371
8.1675 | -34.2612 | -14.1467
8.1736 | -34.2615 | -14.1568
8.1797 | -34.2617 | -14.1668
8.1858 | -34.262 | -14.1765
8.1919 | -34.2623 | -14.1862
8.198 | -34.2626 | -14.1956
8.2041 | -34.2629 | -14.2057
8.2102 | -34.2633 | -14.2155
8.2163 | -34.2636 | -14.2252
8.2224 | -34.2639 | -14.2352
8.2285 | -34.2642 | -14.2447
8.2346 | -34.2645 | -14.2544
8.2408 | -34.2648 | -14.2643
8.2469 | -34.2651 | -14.2739
8.253 | -34.2653 | -14.2836
8.2591 | -34.2657 | -14.2931
8.2652 | -34.2661 | -14.3027
8.2713 | -34.2664 | -14.3125
8.2774 | -34.2667 | -14.3223
8.2835 | -34.267 | -14.3324
8.2896 | -34.2673 | -14.3417
8.2957 | -34.2676 | -14.3514
8.3018 | -34.2679 | -14.3612
8.3079 | -34.2683 | -14.3709
8.314 | -34.2686 | -14.3805
8.3201 | -34.2689 | -14.3902
8.3262 | -34.2692 | -14.3999
8.3323 | -34.2695 | -14.4096
8.3384 | -34.2698 | -14.4197
8.3445 | -34.2701 | -14.429
8.3506 | -34.2704 | -14.439
8.3567 | -34.2707 | -14.449
8.3628 | -34.271 | -14.4587
8.3689 | -34.2713 | -14.4684
8.375 | -34.2716 | -14.4782
8.3812 | -34.2719 | -14.4879
8.3873 | -34.2722 | -14.4977
8.3934 | -34.2725 | -14.5074
8.3995 | -34.2728 | -14.5173
8.4056 | -34.2731 | -14.5271
8.4117 | -34.2734 | -14.5371
8.4178 | -34.2738 | -14.5464
8.4239 | -34.2741 | -14.5559
8.43 | -34.2744 | -14.5657
8.4361 | -34.2747 | -14.5758
8.4422 | -34.275 | -14.5854
8.4483 | -34.2752 | -14.5951
8.4544 | -34.2756 | -14.6052
8.4605 | -34.2759 | -14.6148
8.4666 | -34.2761 | -14.6247
8.4727 | -34.2764 | -14.6344
8.4788 | -34.2767 | -14.6439
8.4849 | -34.277 | -14.654
8.491 | -34.2773 | -14.6636
8.4971 | -34.2776 | -14.6734
8.5032 | -34.278 | -14.6832
8.5093 | -34.2783 | -14.6929
8.5154 | -34.2786 | -14.7024
8.5215 | -34.2788 | -14.7122
8.5277 | -34.2792 | -14.722
8.5338 | -34.2795 | -14.7315
8.5399 | -34.2798 | -14.7415
8.546 | -34.2801 | -14.7509
8.5521 | -34.2804 | -14.7608
8.5582 | -34.2807 | -14.7707
8.5643 | -34.281 | -14.7802
8.5704 | -34.2813 | -14.7896
8.5765 | -34.2816 | -14.7993
8.5826 | -34.282 | -14.809
8.5887 | -34.2823 | -14.8186
8.5948 | -34.2826 | -14.8279
8.6009 | -34.2829 | -14.8376
8.607 | -34.2832 | -14.8475
8.6131 | -34.2835 | -14.8575
8.6192 | -34.2838 | -14.8668
8.6253 | -34.2841 | -14.8767
8.6314 | -34.2844 | -14.8863
8.6375 | -34.2847 | -14.8962
8.6436 | -34.2851 | -14.906
8.6497 | -34.2854 | -14.9158
8.6558 | -34.2857 | -14.9252
8.6619 | -34.286 | -14.9352
8.6681 | -34.2863 | -14.9448
8.6742 | -34.2866 | -14.9547
8.6803 | -34.287 | -14.9647
8.6864 | -34.2873 | -14.9747
8.6925 | -34.2876 | -14.9844
8.6986 | -34.2879 | -14.9945
8.7047 | -34.2883 | -15.0043
8.7108 | -34.2886 | -15.0141
8.7169 | -34.2889 | -15.0236
8.723 | -34.2893 | -15.0333
8.7291 | -34.2895 | -15.0429
8.7352 | -34.2899 | -15.0529
8.7413 | -34.2902 | -15.0626
8.7474 | -34.2905 | -15.072
8.7535 | -34.2909 | -15.0819
8.7596 | -34.2913 | -15.0921
8.7657 | -34.2915 | -15.1018
8.7718 | -34.2918 | -15.1113
8.7779 | -34.2922 | -15.1216
8.784 | -34.2924 | -15.1311
8.7901 | -34.2927 | -15.1413
8.7962 | -34.2931 | -15.1511
8.8023 | -34.2935 | -15.1605
8.8084 | -34.2938 | -15.1704
8.8146 | -34.2941 | -15.18
8.8207 | -34.2944 | -15.1897
8.8268 | -34.2948 | -15.1993
8.8329 | -34.2951 | -15.2092
8.839 | -34.2954 | -15.219
8.8451 | -34.2957 | -15.2286
8.8512 | -34.2961 | -15.2384
8.8573 | -34.2964 | -15.2479
8.8634 | -34.2968 | -15.2576
8.8695 | -34.2971 | -15.2674
8.8756 | -34.2974 | -15.2769
8.8817 | -34.2978 | -15.2869
8.8878 | -34.298 | -15.2969
8.8939 | -34.2983 | -15.3063
8.9 | -34.2987 | -15.3157
8.9061 | -34.299 | -15.3255
8.9122 | -34.2994 | -15.335
8.9183 | -34.2997 | -15.3446
8.9244 | -34.3 | -15.3542
8.9305 | -34.3004 | -15.3639
8.9366 | -34.3007 | -15.3734
8.9427 | -34.3011 | -15.3832
8.9488 | -34.3014 | -15.3929
8.955 | -34.3018 | -15.4026
8.9611 | -34.3021 | -15.4124
8.9672 | -34.3025 | -15.4225
8.9733 | -34.3028 | -15.4321
8.9794 | -34.3032 | -15.4419
8.9855 | -34.3035 | -15.4515
8.9916 | -34.3038 | -15.4611
8.9977 | -34.3042 | -15.4707
9.0038 | -34.3046 | -15.4806
9.0099 | -34.3049 | -15.4904
9.016 | -34.3053 | -15.5001
9.0221 | -34.3056 | -15.5101
9.0282 | -34.306 | -15.5199
9.0343 | -34.3063 | -15.5299
9.0404 | -34.3068 | -15.5395
9.0465 | -34.3071 | -15.5496
9.0526 | -34.3074 | -15.5592
9.0587 | -34.3078 | -15.5688
9.0648 | -34.3082 | -15.5785
9.0709 | -34.3086 | -15.5879
9.077 | -34.3089 | -15.5982
9.0831 | -34.3092 | -15.6079
9.0892 | -34.3095 | -15.6177
9.0953 | -34.3098 | -15.6275
9.1015 | -34.3102 | -15.6372
9.1076 | -34.3106 | -15.647
9.1137 | -34.311 | -15.6566
9.1198 | -34.3113 | -15.6659
9.1259 | -34.3117 | -15.6757
9.132 | -34.3121 | -15.6857
9.1381 | -34.3125 | -15.6953
9.1442 | -34.3129 | -15.7046
9.1503 | -34.3132 | -15.7145
9.1564 | -34.3136 | -15.7243
9.1625 | -34.3139 | -15.734
9.1686 | -34.3143 | -15.7437
9.1747 | -34.3147 | -15.7535
9.1808 | -34.3151 | -15.7634
9.1869 | -34.3154 | -15.7734
9.193 | -34.3158 | -15.7833
9.1991 | -34.3161 | -15.7928
9.2052 | -34.3164 | -15.8024
9.2113 | -34.3169 | -15.8123
9.2174 | -34.3172 | -15.8227
9.2235 | -34.3176 | -15.8321
9.2296 | -34.318 | -15.8419
9.2357 | -34.3184 | -15.8517
9.2419 | -34.3187 | -15.8614
9.248 | -34.3192 | -15.8713
9.2541 | -34.3195 | -15.8812
9.2602 | -34.3199 | -15.891
9.2663 | -34.3202 | -15.9006
9.2724 | -34.3206 | -15.9104
9.2785 | -34.321 | -15.9202
9.2846 | -34.3214 | -15.93
9.2907 | -34.3217 | -15.9396
9.2968 | -34.322 | -15.9494
9.3029 | -34.3224 | -15.9592
9.309 | -34.3228 | -15.9689
9.3151 | -34.3231 | -15.9785
9.3212 | -34.3236 | -15.9884
9.3273 | -34.324 | -15.998
9.3334 | -34.3244 | -16.0075
9.3395 | -34.3249 | -16.0171
9.3456 | -34.3253 | -16.0266
9.3517 | -34.3256 | -16.036
9.3578 | -34.326 | -16.0455
9.3639 | -34.3264 | -16.0555
9.37 | -34.3268 | -16.0653
9.3761 | -34.3272 | -16.0748
9.3822 | -34.3276 | -16.0844
9.3884 | -34.328 | -16.0941
9.3945 | -34.3284 | -16.1039
9.4006 | -34.3287 | -16.1136
9.4067 | -34.3291 | -16.1235
9.4128 | -34.3295 | -16.1331
9.4189 | -34.3299 | -16.143
9.425 | -34.3303 | -16.1526
9.4311 | -34.3307 | -16.1622
9.4372 | -34.3311 | -16.1719
9.4433 | -34.3315 | -16.1818
9.4494 | -34.3318 | -16.1914
9.4555 | -34.3322 | -16.2011
9.4616 | -34.3327 | -16.2106
9.4677 | -34.3331 | -16.2202
9.4738 | -34.3335 | -16.2302
9.4799 | -34.3338 | -16.2398
9.486 | -34.3341 | -16.2498
9.4921 | -34.3346 | -16.2599
9.4982 | -34.335 | -16.2692
9.5043 | -34.3353 | -16.2786
9.5104 | -34.3358 | -16.2881
9.5165 | -34.3362 | -16.2979
9.5226 | -34.3366 | -16.3077
9.5288 | -34.337 | -16.3174
9.5349 | -34.3374 | -16.3271
9.541 | -34.3378 | -16.3367
9.5471 | -34.3381 | -16.3465
9.5532 | -34.3385 | -16.3562
9.5593 | -34.339 | -16.3659
9.5654 | -34.3394 | -16.3756
9.5715 | -34.3398 | -16.3856
9.5776 | -34.3402 | -16.3955
9.5837 | -34.3405 | -16.4049
9.5898 | -34.3409 | -16.4149
9.5959 | -34.3413 | -16.4245
9.602 | -34.3416 | -16.4342
9.6081 | -34.342 | -16.444
9.6142 | -34.3425 | -16.4536
9.6203 | -34.3428 | -16.4633
9.6264 | -34.3432 | -16.4726
9.6325 | -34.3435 | -16.4826
9.6386 | -34.3439 | -16.4922
9.6447 | -34.3443 | -16.5018
9.6508 | -34.3447 | -16.5116
9.6569 | -34.3451 | -16.5212
9.663 | -34.3455 | -16.531
9.6691 | -34.3459 | -16.5407
9.6753 | -34.3464 | -16.5503
9.6814 | -34.3468 | -16.5598
9.6875 | -34.3472 | -16.5693
9.6936 | -34.3476 | -16.5788
9.6997 | -34.348 | -16.5887
9.7058 | -34.3485 | -16.5983
9.7119 | -34.3488 | -16.6079
9.718 | -34.3492 | -16.6179
9.7241 | -34.3496 | -16.6276
9.7302 | -34.35 | -16.6373
9.7363 | -34.3504 | -16.6467
9.7424 | -34.3508 | -16.6562
9.7485 | -34.3512 | -16.6658
9.7546 | -34.3516 | -16.6752
9.7607 | -34.352 | -16.6851
9.7668 | -34.3524 | -16.6951
9.7729 | -34.3529 | -16.7045
9.779 | -34.3533 | -16.7141
9.7851 | -34.3537 | -16.7239
9.7912 | -34.3541 | -16.7341
9.7973 | -34.3544 | -16.7439
9.8034 | -34.3549 | -16.7538
9.8095 | -34.3553 | -16.7633
9.8157 | -34.3556 | -16.773
9.8218 | -34.356 | -16.7823
9.8279 | -34.3565 | -16.7921
9.834 | -34.3569 | -16.8016
9.8401 | -34.3573 | -16.8112
9.8462 | -34.3578 | -16.8205
9.8523 | -34.3582 | -16.8301
9.8584 | -34.3586 | -16.8399
9.8645 | -34.3589 | -16.8497
9.8706 | -34.3593 | -16.8592
9.8767 | -34.3597 | -16.8684
9.8828 | -34.3602 | -16.8777
9.8889 | -34.3606 | -16.8875
9.895 | -34.361 | -16.8969
9.9011 | -34.3613 | -16.906
9.9072 | -34.3618 | -16.9152
9.9133 | -34.3622 | -16.9246
9.9194 | -34.3626 | -16.934
9.9255 | -34.363 | -16.9437
9.9316 | -34.3635 | -16.953
9.9377 | -34.3639 | -16.9625
9.9438 | -34.3643 | -16.972
9.9499 | -34.3647 | -16.9817
9.956 | -34.3651 | -16.9911
9.9622 | -34.3655 | -17.0005
9.9683 | -34.3659 | -17.0102
9.9744 | -34.3663 | -17.0197
9.9805 | -34.3668 | -17.0293
9.9866 | -34.3672 | -17.0386
9.9927 | -34.3676 | -17.0483
9.9988 | -34.3681 | -17.0577
10.0049 | -34.3684 | -17.0672
10.011 | -34.3689 | -17.0766
10.0171 | -34.3693 | -17.0861
10.0232 | -34.3697 | -17.0954
10.0293 | -34.3702 | -17.1047
10.0354 | -34.3707 | -17.1139
10.0415 | -34.3711 | -17.1234
10.0476 | -34.3715 | -17.1332
10.0537 | -34.3718 | -17.143
10.0598 | -34.3723 | -17.1526
10.0659 | -34.3727 | -17.1622
10.072 | -34.3731 | -17.1715
10.0781 | -34.3736 | -17.181
10.0842 | -34.374 | -17.1906
10.0903 | -34.3744 | -17.2004
10.0964 | -34.3749 | -17.2097
10.1026 | -34.3753 | -17.2192
10.1087 | -34.3758 | -17.2288
10.1148 | -34.3762 | -17.2384
10.1209 | -34.3766 | -17.2477
10.127 | -34.377 | -17.257
10.1331 | -34.3775 | -17.2664
10.1392 | -34.3779 | -17.2758
10.1453 | -34.3783 | -17.2853
10.1514 | -34.3787 | -17.2945
10.1575 | -34.3791 | -17.3039
10.1636 | -34.3796 | -17.3127
10.1697 | -34.38 | -17.3221
10.1758 | -34.3805 | -17.3314
10.1819 | -34.3809 | -17.3409
10.188 | -34.3813 | -17.3501
10.1941 | -34.3817 | -17.3596
10.2002 | -34.382 | -17.3684
10.2063 | -34.3825 | -17.3781
10.2124 | -34.3829 | -17.3876
10.2185 | -34.3833 | -17.3969
10.2246 | -34.3838 | -17.4063
10.2307 | -34.3841 | -17.4158
10.2368 | -34.3846 | -17.4244
10.2429 | -34.3851 | -17.4335
10.2491 | -34.3855 | -17.4427
10.2552 | -34.3859 | -17.4525
10.2613 | -34.3864 | -17.4622
10.2674 | -34.3867 | -17.4718
10.2735 | -34.3871 | -17.4814
10.2796 | -34.3875 | -17.4909
10.2857 | -34.3879 | -17.5003
10.2918 | -34.3883 | -17.5098
10.2979 | -34.3888 | -17.5194
10.304 | -34.3893 | -17.529
10.3101 | -34.3897 | -17.5382
10.3162 | -34.3901 | -17.5475
10.3223 | -34.3906 | -17.5568
10.3284 | -34.391 | -17.5661
10.3345 | -34.3914 | -17.5753
10.3406 | -34.3918 | -17.5844
10.3467 | -34.3922 | -17.5937
10.3528 | -34.3926 | -17.6029
10.3589 | -34.393 | -17.6124
10.365 | -34.3935 | -17.622
10.3711 | -34.3939 | -17.6316
10.3772 | -34.3942 | -17.6412
10.3833 | -34.3947 | -17.6505
10.3895 | -34.3951 | -17.6596
10.3956 | -34.3955 | -17.6687
10.4017 | -34.3959 | -17.6781
10.4078 | -34.3963 | -17.6875
10.4139 | -34.3968 | -17.6969
10.42 | -34.3972 | -17.7063
10.4261 | -34.3977 | -17.7156
10.4322 | -34.3981 | -17.7248
10.4383 | -34.3985 | -17.7342
10.4444 | -34.3989 | -17.7432
10.4505 | -34.3994 | -17.7528
10.4566 | -34.3998 | -17.762
10.4627 | -34.4002 | -17.7713
10.4688 | -34.4006 | -17.7801
10.4749 | -34.4011 | -17.7895
10.481 | -34.4015 | -17.7989
10.4871 | -34.402 | -17.8084
10.4932 | -34.4024 | -17.8175
10.4993 | -34.4029 | -17.8267
10.5054 | -34.4033 | -17.8359
10.5115 | -34.4038 | -17.8453
10.5176 | -34.4042 | -17.8543
10.5237 | -34.4047 | -17.8636
10.5298 | -34.405 | -17.8727
10.536 | -34.4055 | -17.8821
10.5421 | -34.4059 | -17.8908
10.5482 | -34.4063 | -17.9004
10.5543 | -34.4069 | -17.9098
10.5604 | -34.4073 | -17.9187
10.5665 | -34.4077 | -17.9278
10.5726 | -34.4081 | -17.9369
10.5787 | -34.4086 | -17.9463
10.5848 | -34.4091 | -17.9552
10.5909 | -34.4095 | -17.9646
10.597 | -34.4099 | -17.9738
10.6031 | -34.4103 | -17.9833
10.6092 | -34.4108 | -17.9928
10.6153 | -34.4111 | -18.0023
10.6214 | -34.4116 | -18.0117
10.6275 | -34.4121 | -18.0206
10.6336 | -34.4125 | -18.0301
10.6397 | -34.4129 | -18.0393
10.6458 | -34.4133 | -18.0483
10.6519 | -34.4137 | -18.0576
10.658 | -34.4141 | -18.0666
10.6641 | -34.4146 | -18.0759
10.6702 | -34.415 | -18.0854
10.6764 | -34.4155 | -18.0946
10.6825 | -34.4159 | -18.104
10.6886 | -34.4164 | -18.1133
10.6947 | -34.4168 | -18.1223
10.7008 | -34.4173 | -18.1315
10.7069 | -34.4177 | -18.1408
10.713 | -34.4182 | -18.1501
10.7191 | -34.4186 | -18.1596
10.7252 | -34.4191 | -18.1687
10.7313 | -34.4195 | -18.1779
10.7374 | -34.4199 | -18.187
10.7435 | -34.4203 | -18.1965
10.7496 | -34.4208 | -18.2055
10.7557 | -34.4212 | -18.2149
10.7618 | -34.4217 | -18.2242
10.7679 | -34.4221 | -18.2333
10.774 | -34.4225 | -18.2425
10.7801 | -34.4229 | -18.2521
10.7862 | -34.4234 | -18.261
10.7923 | -34.4239 | -18.2702
10.7984 | -34.4243 | -18.2795
10.8045 | -34.4248 | -18.2889
10.8106 | -34.4252 | -18.2982
10.8168 | -34.4257 | -18.3076
10.8229 | -34.4261 | -18.3169
10.829 | -34.4265 | -18.3263
10.8351 | -34.4269 | -18.3356
10.8412 | -34.4273 | -18.3446
10.8473 | -34.4277 | -18.3535
10.8534 | -34.4282 | -18.3629
10.8595 | -34.4286 | -18.372
10.8656 | -34.4291 | -18.3812
10.8717 | -34.4296 | -18.3901
10.8778 | -34.43 | -18.3991
10.8839 | -34.4304 | -18.4082
10.89 | -34.4309 | -18.4169
10.8961 | -34.4313 | -18.4261
10.9022 | -34.4317 | -18.4354
10.9083 | -34.4322 | -18.4449
10.9144 | -34.4326 | -18.4539
10.9205 | -34.4331 | -18.4635
10.9266 | -34.4336 | -18.4729
10.9327 | -34.434 | -18.4822
10.9388 | -34.4344 | -18.4918
10.9449 | -34.4348 | -18.5011
10.951 | -34.4353 | -18.5103
10.9571 | -34.4357 | -18.5197
10.9633 | -34.4362 | -18.529
10.9694 | -34.4366 | -18.5385
10.9755 | -34.4372 | -18.5475
10.9816 | -34.4376 | -18.5566
10.9877 | -34.438 | -18.5659
10.9938 | -34.4385 | -18.5757
10.9999 | -34.4389 | -18.585
11.006 | -34.4394 | -18.5942
11.0121 | -34.4399 | -18.6039
11.0182 | -34.4403 | -18.6129
11.0243 | -34.4408 | -18.622
11.0304 | -34.4412 | -18.6314
11.0365 | -34.4417 | -18.6405
11.0426 | -34.4421 | -18.6495
11.0487 | -34.4426 | -18.6587
11.0548 | -34.4431 | -18.6678
11.0609 | -34.4435 | -18.6773
11.067 | -34.4439 | -18.6868
11.0731 | -34.4444 | -18.6961
11.0792 | -34.4449 | -18.7051
11.0853 | -34.4453 | -18.7141
11.0914 | -34.4457 | -18.7231
11.0975 | -34.4462 | -18.7324
11.1037 | -34.4467 | -18.7412
11.1098 | -34.4471 | -18.7503
11.1159 | -34.4476 | -18.7597
11.122 | -34.448 | -18.7692
11.1281 | -34.4485 | -18.7781
11.1342 | -34.4489 | -18.7868
11.1403 | -34.4494 | -18.7963
11.1464 | -34.4499 | -18.8055
11.1525 | -34.4503 | -18.8147
11.1586 | -34.4507 | -18.8247
11.1647 | -34.4512 | -18.8337
11.1708 | -34.4516 | -18.8425
11.1769 | -34.452 | -18.851
11.183 | -34.4525 | -18.8599
11.1891 | -34.453 | -18.8693
11.1952 | -34.4534 | -18.8788
11.2013 | -34.4539 | -18.8881
11.2074 | -34.4543 | -18.8972
11.2135 | -34.4547 | -18.9065
11.2196 | -34.4551 | -18.9152
11.2257 | -34.4555 | -18.9241
11.2318 | -34.4559 | -18.9337
11.2379 | -34.4563 | -18.9426
11.244 | -34.4567 | -18.9517
11.2502 | -34.4572 | -18.9611
11.2563 | -34.4577 | -18.9707
11.2624 | -34.4581 | -18.98
11.2685 | -34.4586 | -18.9889
11.2746 | -34.459 | -18.9979
11.2807 | -34.4594 | -19.0077
11.2868 | -34.4598 | -19.0173
11.2929 | -34.4603 | -19.0261
11.299 | -34.4608 | -19.0355
11.3051 | -34.4612 | -19.0445
11.3112 | -34.4616 | -19.0534
11.3173 | -34.4621 | -19.0622
11.3234 | -34.4624 | -19.0714
11.3295 | -34.4629 | -19.0805
11.3356 | -34.4634 | -19.0895
11.3417 | -34.4638 | -19.099
11.3478 | -34.4641 | -19.1083
11.3539 | -34.4646 | -19.1173
11.36 | -34.4649 | -19.1263
11.3661 | -34.4654 | -19.1355
11.3722 | -34.4658 | -19.1444
11.3783 | -34.4663 | -19.1537
11.3844 | -34.4667 | -19.1632
11.3906 | -34.4671 | -19.1728
11.3967 | -34.4674 | -19.1818
11.4028 | -34.4679 | -19.191
11.4089 | -34.4683 | -19.2001
11.415 | -34.4687 | -19.2092
11.4211 | -34.4692 | -19.2184
11.4272 | -34.4696 | -19.2277
11.4333 | -34.47 | -19.2372
11.4394 | -34.4704 | -19.2466
11.4455 | -34.4708 | -19.2557
11.4516 | -34.4713 | -19.2648
11.4577 | -34.4717 | -19.2741
11.4638 | -34.4722 | -19.2828
11.4699 | -34.4726 | -19.2922
11.476 | -34.473 | -19.3016
11.4821 | -34.4734 | -19.3102
11.4882 | -34.4739 | -19.3197
11.4943 | -34.4743 | -19.3286
11.5004 | -34.4749 | -19.338
11.5065 | -34.4753 | -19.347
11.5126 | -34.4758 | -19.3559
11.5187 | -34.4762 | -19.3649
11.5248 | -34.4767 | -19.3739
11.5309 | -34.4772 | -19.3831
11.5371 | -34.4776 | -19.3925
11.5432 | -34.478 | -19.4018
11.5493 | -34.4784 | -19.411
11.5554 | -34.4789 | -19.4199
11.5615 | -34.4793 | -19.4292
11.5676 | -34.4798 | -19.4385
11.5737 | -34.4803 | -19.4473
11.5798 | -34.4807 | -19.4564
11.5859 | -34.4812 | -19.4654
11.592 | -34.4816 | -19.4747
11.5981 | -34.482 | -19.4837
11.6042 | -34.4824 | -19.493
11.6103 | -34.4828 | -19.502
11.6164 | -34.4833 | -19.5108
11.6225 | -34.4837 | -19.5199
11.6286 | -34.4842 | -19.529
11.6347 | -34.4846 | -19.5385
11.6408 | -34.4851 | -19.5477
11.6469 | -34.4855 | -19.5572
11.653 | -34.486 | -19.5661
11.6591 | -34.4864 | -19.5753
11.6652 | -34.4869 | -19.5845
11.6713 | -34.4873 | -19.5934
11.6775 | -34.4877 | -19.6027
11.6836 | -34.4881 | -19.6119
11.6897 | -34.4886 | -19.6208
11.6958 | -34.4891 | -19.63
11.7019 | -34.4895 | -19.6391
11.708 | -34.49 | -19.6484
11.7141 | -34.4905 | -19.6576
11.7202 | -34.4909 | -19.6663
11.7263 | -34.4913 | -19.6756
11.7324 | -34.4918 | -19.6848
11.7385 | -34.4922 | -19.6941
11.7446 | -34.4926 | -19.7033
11.7507 | -34.4931 | -19.7122
11.7568 | -34.4935 | -19.7216
11.7629 | -34.4939 | -19.7306
11.769 | -34.4944 | -19.7397
11.7751 | -34.4948 | -19.7485
11.7812 | -34.4952 | -19.7573
11.7873 | -34.4957 | -19.7667
11.7934 | -34.4962 | -19.7758
11.7995 | -34.4966 | -19.7845
11.8056 | -34.497 | -19.7939
11.8117 | -34.4974 | -19.803
11.8178 | -34.4978 | -19.8121
11.824 | -34.4983 | -19.8209
11.8301 | -34.4988 | -19.8297
11.8362 | -34.4992 | -19.8384
11.8423 | -34.4996 | -19.8476
11.8484 | -34.5 | -19.8567
11.8545 | -34.5005 | -19.8662
11.8606 | -34.501 | -19.8753
11.8667 | -34.5014 | -19.8845
11.8728 | -34.5018 | -19.8937
11.8789 | -34.5023 | -19.9032
11.885 | -34.5026 | -19.9122
11.8911 | -34.5031 | -19.9214
11.8972 | -34.5036 | -19.9301
11.9033 | -34.504 | -19.9388
11.9094 | -34.5044 | -19.9476
11.9155 | -34.5048 | -19.9565
11.9216 | -34.5053 | -19.9655
11.9277 | -34.5057 | -19.9742
11.9338 | -34.5061 | -19.9834
11.9399 | -34.5065 | -19.9928
11.946 | -34.507 | -20.0019
11.9521 | -34.5074 | -20.0107
11.9582 | -34.5078 | -20.0197
11.9644 | -34.5083 | -20.0286
11.9705 | -34.5087 | -20.0378
11.9766 | -34.5092 | -20.0467
11.9827 | -34.5096 | -20.0557
11.9888 | -34.51 | -20.0645
11.9949 | -34.5105 | -20.0734
12.001 | -34.5109 | -20.0825
12.0071 | -34.5114 | -20.0914
12.0132 | -34.5118 | -20.1003
12.0193 | -34.5123 | -20.1095
12.0254 | -34.5127 | -20.1184
12.0315 | -34.5132 | -20.1272
12.0376 | -34.5136 | -20.1358
12.0437 | -34.514 | -20.1446
12.0498 | -34.5144 | -20.1539
12.0559 | -34.5149 | -20.1629
12.062 | -34.5153 | -20.1717
12.0681 | -34.5158 | -20.1806
12.0742 | -34.5162 | -20.1898
12.0803 | -34.5167 | -20.1986
12.0864 | -34.5171 | -20.2076
12.0925 | -34.5176 | -20.2166
12.0986 | -34.518 | -20.2255
12.1047 | -34.5185 | -20.2344
12.1109 | -34.5189 | -20.2435
12.117 | -34.5194 | -20.2523
12.1231 | -34.5199 | -20.2613
12.1292 | -34.5203 | -20.2702
12.1353 | -34.5207 | -20.2789
12.1414 | -34.5212 | -20.2881
12.1475 | -34.5216 | -20.297
12.1536 | -34.522 | -20.3059
12.1597 | -34.5225 | -20.3151
12.1658 | -34.5229 | -20.324
12.1719 | -34.5234 | -20.3331
12.178 | -34.5238 | -20.3422
12.1841 | -34.5242 | -20.3513
12.1902 | -34.5247 | -20.3603
12.1963 | -34.5251 | -20.3692
12.2024 | -34.5256 | -20.3779
12.2085 | -34.526 | -20.3868
12.2146 | -34.5264 | -20.3955
12.2207 | -34.5269 | -20.4045
12.2268 | -34.5273 | -20.4134
12.2329 | -34.5277 | -20.4223
12.239 | -34.5281 | -20.4308
12.2451 | -34.5286 | -20.4394
12.2513 | -34.5291 | -20.4485
12.2574 | -34.5295 | -20.4576
12.2635 | -34.53 | -20.4662
12.2696 | -34.5304 | -20.4749
12.2757 | -34.5308 | -20.4835
12.2818 | -34.5312 | -20.4921
12.2879 | -34.5316 | -20.5008
12.294 | -34.5321 | -20.5096
12.3001 | -34.5325 | -20.5187
12.3062 | -34.533 | -20.5276
12.3123 | -34.5334 | -20.5365
12.3184 | -34.5338 | -20.5451
12.3245 | -34.5343 | -20.554
12.3306 | -34.5347 | -20.5625
12.3367 | -34.5352 | -20.5712
12.3428 | -34.5356 | -20.5798
12.3489 | -34.5361 | -20.5881
12.355 | -34.5365 | -20.5969
12.3611 | -34.5369 | -20.6056
12.3672 | -34.5373 | -20.6142
12.3733 | -34.5378 | -20.6234
12.3794 | -34.5382 | -20.6317
12.3855 | -34.5386 | -20.6403
12.3916 | -34.539 | -20.6493
12.3978 | -34.5395 | -20.6584
12.4039 | -34.5399 | -20.6671
12.41 | -34.5403 | -20.676
12.4161 | -34.5408 | -20.6844
12.4222 | -34.5412 | -20.6931
12.4283 | -34.5416 | -20.7018
12.4344 | -34.5421 | -20.7106
12.4405 | -34.5425 | -20.7195
12.4466 | -34.5429 | -20.728
12.4527 | -34.5434 | -20.7367
12.4588 | -34.5438 | -20.7453
12.4649 | -34.5443 | -20.7543
12.471 | -34.5447 | -20.7635
12.4771 | -34.5451 | -20.7718
12.4832 | -34.5456 | -20.7806
12.4893 | -34.546 | -20.7891
12.4954 | -34.5465 | -20.7977
12.5015 | -34.547 | -20.8062
12.5076 | -34.5474 | -20.8149
12.5137 | -34.5478 | -20.824
12.5198 | -34.5482 | -20.8323
12.5259 | -34.5487 | -20.841
12.532 | -34.5491 | -20.8499
12.5382 | -34.5496 | -20.8586
12.5443 | -34.55 | -20.8674
12.5504 | -34.5504 | -20.8762
12.5565 | -34.5509 | -20.8846
12.5626 | -34.5514 | -20.8935
12.5687 | -34.5518 | -20.9024
12.5748 | -34.5522 | -20.911
12.5809 | -34.5526 | -20.9193
12.587 | -34.5531 | -20.9284
12.5931 | -34.5535 | -20.937
12.5992 | -34.554 | -20.946
12.6053 | -34.5544 | -20.9544
12.6114 | -34.5549 | -20.963
12.6175 | -34.5553 | -20.9718
12.6236 | -34.5558 | -20.9802
12.6297 | -34.5563 | -20.9886
12.6358 | -34.5567 | -20.997
12.6419 | -34.5572 | -21.006
12.648 | -34.5576 | -21.0149
12.6541 | -34.5581 | -21.0235
12.6602 | -34.5585 | -21.0316
12.6663 | -34.5589 | -21.04
12.6724 | -34.5594 | -21.0488
12.6785 | -34.5598 | -21.0577
12.6847 | -34.5602 | -21.0664
12.6908 | -34.5607 | -21.0751
12.6969 | -34.5612 | -21.0836
12.703 | -34.5616 | -21.0922
12.7091 | -34.562 | -21.101
12.7152 | -34.5624 | -21.1096
12.7213 | -34.5629 | -21.1183
12.7274 | -34.5633 | -21.1271
12.7335 | -34.5637 | -21.1349
12.7396 | -34.5641 | -21.1431
12.7457 | -34.5646 | -21.1517
12.7518 | -34.565 | -21.1603
12.7579 | -34.5654 | -21.1693
12.764 | -34.5658 | -21.178
12.7701 | -34.5662 | -21.1862
12.7762 | -34.5666 | -21.1948
12.7823 | -34.5671 | -21.2035
12.7884 | -34.5674 | -21.2121
12.7945 | -34.5679 | -21.2203
12.8006 | -34.5683 | -21.2282
12.8067 | -34.5687 | -21.237
12.8128 | -34.5692 | -21.2452
12.8189 | -34.5695 | -21.2535
12.8251 | -34.5699 | -21.2618
12.8312 | -34.5703 | -21.2704
12.8373 | -34.5708 | -21.2789
12.8434 | -34.5712 | -21.2874
12.8495 | -34.5716 | -21.2959
12.8556 | -34.572 | -21.3044
12.8617 | -34.5725 | -21.3127
12.8678 | -34.5729 | -21.3213
12.8739 | -34.5733 | -21.3294
12.88 | -34.5737 | -21.3381
12.8861 | -34.5741 | -21.3466
12.8922 | -34.5745 | -21.355
12.8983 | -34.575 | -21.3632
12.9044 | -34.5755 | -21.3715
12.9105 | -34.576 | -21.3799
12.9166 | -34.5764 | -21.3887
12.9227 | -34.5768 | -21.3976
12.9288 | -34.5773 | -21.4063
12.9349 | -34.5777 | -21.4151
12.941 | -34.5782 | -21.4237
12.9471 | -34.5787 | -21.4322
12.9532 | -34.5792 | -21.4407
12.9593 | -34.5796 | -21.4494
12.9654 | -34.58 | -21.4582
12.9716 | -34.5804 | -21.4668
12.9777 | -34.5808 | -21.4751
12.9838 | -34.5812 | -21.4835
12.9899 | -34.5816 | -21.492
12.996 | -34.582 | -21.5006
13.0021 | -34.5824 | -21.5092
13.0082 | -34.5829 | -21.5178
13.0143 | -34.5833 | -21.5258
13.0204 | -34.5836 | -21.5339
13.0265 | -34.5841 | -21.5421
13.0326 | -34.5845 | -21.5507
13.0387 | -34.585 | -21.5595
13.0448 | -34.5854 | -21.5679
13.0509 | -34.5858 | -21.5766
13.057 | -34.5863 | -21.5847
13.0631 | -34.5868 | -21.593
13.0692 | -34.5872 | -21.6012
13.0753 | -34.5877 | -21.6097
13.0814 | -34.5881 | -21.6179
13.0875 | -34.5885 | -21.6262
13.0936 | -34.589 | -21.6344
13.0997 | -34.5894 | -21.6424
13.1058 | -34.5898 | -21.6504
13.112 | -34.5901 | -21.6588
13.1181 | -34.5906 | -21.6672
13.1242 | -34.591 | -21.6757
13.1303 | -34.5915 | -21.6843
13.1364 | -34.5919 | -21.6927
13.1425 | -34.5923 | -21.7011
13.1486 | -34.5927 | -21.7094
13.1547 | -34.5932 | -21.7176
13.1608 | -34.5936 | -21.7259
13.1669 | -34.594 | -21.7343
13.173 | -34.5944 | -21.7426
13.1791 | -34.5949 | -21.751
13.1852 | -34.5953 | -21.7597
13.1913 | -34.5957 | -21.7682
13.1974 | -34.5961 | -21.7766
13.2035 | -34.5965 | -21.785
13.2096 | -34.597 | -21.7934
13.2157 | -34.5975 | -21.8018
13.2218 | -34.5979 | -21.8103
13.2279 | -34.5983 | -21.8186
13.234 | -34.5987 | -21.827
13.2401 | -34.5992 | -21.8354
13.2462 | -34.5996 | -21.8436
13.2524 | -34.6 | -21.8518
13.2585 | -34.6004 | -21.8602
13.2646 | -34.6008 | -21.8686
13.2707 | -34.6013 | -21.8771
13.2768 | -34.6017 | -21.8858
13.2829 | -34.602 | -21.8942
13.289 | -34.6025 | -21.9026
13.2951 | -34.6029 | -21.9107
13.3012 | -34.6034 | -21.9191
13.3073 | -34.6038 | -21.9278
13.3134 | -34.6042 | -21.9361
13.3195 | -34.6046 | -21.9443
13.3256 | -34.6051 | -21.9529
13.3317 | -34.6055 | -21.9611
13.3378 | -34.6059 | -21.9695
13.3439 | -34.6063 | -21.9778
13.35 | -34.6067 | -21.9862
13.3561 | -34.6071 | -21.9947
13.3622 | -34.6075 | -22.003
13.3683 | -34.608 | -22.0111
13.3744 | -34.6084 | -22.0195
13.3805 | -34.6087 | -22.028
13.3866 | -34.6091 | -22.0361
13.3927 | -34.6095 | -22.0444
13.3989 | -34.6099 | -22.0525
13.405 | -34.6104 | -22.0607
13.4111 | -34.6109 | -22.0688
13.4172 | -34.6113 | -22.0768
13.4233 | -34.6116 | -22.0855
13.4294 | -34.6121 | -22.0934
13.4355 | -34.6125 | -22.1025
13.4416 | -34.6129 | -22.1109
13.4477 | -34.6132 | -22.1193
13.4538 | -34.6137 | -22.1277
13.4599 | -34.6141 | -22.1363
13.466 | -34.6145 | -22.145
13.4721 | -34.6149 | -22.1535
13.4782 | -34.6154 | -22.1617
13.4843 | -34.6158 | -22.1704
13.4904 | -34.6162 | -22.1782
13.4965 | -34.6166 | -22.1868
13.5026 | -34.6171 | -22.1952
13.5087 | -34.6175 | -22.2035
13.5148 | -34.6179 | -22.2118
13.5209 | -34.6183 | -22.2201
13.527 | -34.6187 | -22.2288
13.5331 | -34.6191 | -22.237
13.5393 | -34.6196 | -22.2453
13.5454 | -34.6199 | -22.2539
13.5515 | -34.6204 | -22.2622
13.5576 | -34.6208 | -22.2706
13.5637 | -34.6213 | -22.2788
13.5698 | -34.6217 | -22.2875
13.5759 | -34.6221 | -22.2953
13.582 | -34.6224 | -22.3038
13.5881 | -34.6229 | -22.3121
13.5942 | -34.6233 | -22.3205
13.6003 | -34.6237 | -22.3291
13.6064 | -34.6241 | -22.3377
13.6125 | -34.6246 | -22.3462
13.6186 | -34.625 | -22.3545
13.6247 | -34.6254 | -22.3629
13.6308 | -34.6258 | -22.3709
13.6369 | -34.6262 | -22.3791
13.643 | -34.6267 | -22.3873
13.6491 | -34.6271 | -22.3957
13.6552 | -34.6275 | -22.4039
13.6613 | -34.628 | -22.4124
13.6674 | -34.6284 | -22.4209
13.6735 | -34.6288 | -22.4294
13.6796 | -34.6292 | -22.4376
13.6858 | -34.6296 | -22.4461
13.6919 | -34.63 | -22.4547
13.698 | -34.6304 | -22.463
13.7041 | -34.6307 | -22.4716
13.7102 | -34.6312 | -22.4802
13.7163 | -34.6317 | -22.4885
13.7224 | -34.6321 | -22.4967
13.7285 | -34.6325 | -22.505
13.7346 | -34.6329 | -22.5134
13.7407 | -34.6333 | -22.5217
13.7468 | -34.6338 | -22.5302
13.7529 | -34.6342 | -22.5383
13.759 | -34.6346 | -22.5468
13.7651 | -34.635 | -22.5553
13.7712 | -34.6354 | -22.5638
13.7773 | -34.6358 | -22.5724
13.7834 | -34.6362 | -22.5807
13.7895 | -34.6367 | -22.5889
13.7956 | -34.6371 | -22.5976
13.8017 | -34.6376 | -22.6057
13.8078 | -34.638 | -22.6141
13.8139 | -34.6384 | -22.6222
13.82 | -34.6388 | -22.6304
13.8262 | -34.6393 | -22.6391
13.8323 | -34.6397 | -22.6476
13.8384 | -34.6401 | -22.6562
13.8445 | -34.6405 | -22.6648
13.8506 | -34.6409 | -22.6732
13.8567 | -34.6414 | -22.6818
13.8628 | -34.6418 | -22.6903
13.8689 | -34.6422 | -22.6988
13.875 | -34.6426 | -22.7071
13.8811 | -34.6431 | -22.7154
13.8872 | -34.6435 | -22.7236
13.8933 | -34.644 | -22.7318
13.8994 | -34.6444 | -22.74
13.9055 | -34.6448 | -22.7482
13.9116 | -34.6453 | -22.7563
13.9177 | -34.6457 | -22.7643
13.9238 | -34.6462 | -22.7729
13.9299 | -34.6467 | -22.7814
13.936 | -34.647 | -22.7897
13.9421 | -34.6475 | -22.798
13.9482 | -34.6479 | -22.8064
13.9543 | -34.6483 | -22.8145
13.9604 | -34.6488 | -22.8231
13.9665 | -34.6493 | -22.8315
13.9727 | -34.6497 | -22.8396
13.9788 | -34.6501 | -22.8482
13.9849 | -34.6504 | -22.8568
13.991 | -34.6508 | -22.8655
13.9971 | -34.6513 | -22.8742
14.0032 | -34.6516 | -22.8828
14.0093 | -34.652 | -22.8911
14.0154 | -34.6524 | -22.8998
14.0215 | -34.6528 | -22.9079
14.0276 | -34.6533 | -22.9163
14.0337 | -34.6538 | -22.925
14.0398 | -34.6542 | -22.9332
14.0459 | -34.6545 | -22.9418
14.052 | -34.6549 | -22.9501
14.0581 | -34.6554 | -22.9581
14.0642 | -34.6558 | -22.9667
14.0703 | -34.6562 | -22.975
14.0764 | -34.6566 | -22.9837
14.0825 | -34.657 | -22.9922
14.0886 | -34.6575 | -23.0001
14.0947 | -34.6578 | -23.0082
14.1008 | -34.6583 | -23.0168
14.1069 | -34.6587 | -23.0252
14.1131 | -34.6591 | -23.0337
14.1192 | -34.6595 | -23.0417
14.1253 | -34.6599 | -23.0499
14.1314 | -34.6604 | -23.0582
14.1375 | -34.6608 | -23.0667
14.1436 | -34.6613 | -23.075
14.1497 | -34.6617 | -23.0831
14.1558 | -34.6621 | -23.0917
14.1619 | -34.6626 | -23.1002
14.168 | -34.663 | -23.1082
14.1741 | -34.6635 | -23.1166
14.1802 | -34.6639 | -23.1248
14.1863 | -34.6643 | -23.1333
14.1924 | -34.6647 | -23.1417
14.1985 | -34.6651 | -23.1499
14.2046 | -34.6656 | -23.158
14.2107 | -34.666 | -23.1659
14.2168 | -34.6664 | -23.1748
14.2229 | -34.6669 | -23.1832
14.229 | -34.6673 | -23.1915
14.2351 | -34.6678 | -23.1996
14.2412 | -34.6682 | -23.2078
14.2473 | -34.6687 | -23.2162
14.2534 | -34.6691 | -23.2245
14.2596 | -34.6695 | -23.2328
14.2657 | -34.67 | -23.2413
14.2718 | -34.6704 | -23.2497
14.2779 | -34.6708 | -23.2579
14.284 | -34.6712 | -23.2663
14.2901 | -34.6717 | -23.2744
14.2962 | -34.6722 | -23.2824
14.3023 | -34.6726 | -23.2909
14.3084 | -34.673 | -23.2994
14.3145 | -34.6735 | -23.3076
14.3206 | -34.6739 | -23.316
14.3267 | -34.6743 | -23.3242
14.3328 | -34.6748 | -23.3325
14.3389 | -34.6752 | -23.3407
14.345 | -34.6756 | -23.3491
14.3511 | -34.676 | -23.3571
14.3572 | -34.6765 | -23.3655
14.3633 | -34.6769 | -23.3738
14.3694 | -34.6773 | -23.382
14.3755 | -34.6777 | -23.3898
14.3816 | -34.6782 | -23.3981
14.3877 | -34.6786 | -23.4062
14.3938 | -34.6791 | -23.4148
14.4 | -34.6795 | -23.423
14.4061 | -34.6799 | -23.4312
14.4122 | -34.6802 | -23.4397
14.4183 | -34.6806 | -23.448
14.4244 | -34.6811 | -23.4561
14.4305 | -34.6815 | -23.4647
14.4366 | -34.682 | -23.4731
14.4427 | -34.6824 | -23.4816
14.4488 | -34.6828 | -23.4897
14.4549 | -34.6833 | -23.498
14.461 | -34.6837 | -23.5063
14.4671 | -34.6842 | -23.5146
14.4732 | -34.6846 | -23.5226
14.4793 | -34.685 | -23.5317
14.4854 | -34.6855 | -23.5396
14.4915 | -34.6859 | -23.5476
14.4976 | -34.6864 | -23.556
14.5037 | -34.6869 | -23.5644
14.5098 | -34.6873 | -23.5726
14.5159 | -34.6878 | -23.581
14.522 | -34.6882 | -23.5894
14.5281 | -34.6887 | -23.5975
14.5342 | -34.6891 | -23.6059
14.5403 | -34.6895 | -23.6143
14.5465 | -34.69 | -23.623
14.5526 | -34.6904 | -23.6315
14.5587 | -34.6908 | -23.6398
14.5648 | -34.6913 | -23.6478
14.5709 | -34.6917 | -23.6562
14.577 | -34.6922 | -23.6645
14.5831 | -34.6926 | -23.6729
14.5892 | -34.6931 | -23.6809
14.5953 | -34.6936 | -23.6889
14.6014 | -34.694 | -23.6971
14.6075 | -34.6945 | -23.7051
14.6136 | -34.6949 | -23.7131
14.6197 | -34.6953 | -23.7215
14.6258 | -34.6958 | -23.7298
14.6319 | -34.6962 | -23.7378
14.638 | -34.6967 | -23.746
14.6441 | -34.6971 | -23.7539
14.6502 | -34.6975 | -23.7619
14.6563 | -34.698 | -23.7702
14.6624 | -34.6984 | -23.7786
14.6685 | -34.6989 | -23.7865
14.6746 | -34.6993 | -23.7947
14.6807 | -34.6998 | -23.8031
14.6869 | -34.7002 | -23.8113
14.693 | -34.7006 | -23.8194
14.6991 | -34.701 | -23.8275
14.7052 | -34.7016 | -23.8356
14.7113 | -34.702 | -23.8441
14.7174 | -34.7024 | -23.8522
14.7235 | -34.7029 | -23.8602
14.7296 | -34.7033 | -23.8683
14.7357 | -34.7038 | -23.8763
14.7418 | -34.7042 | -23.884
14.7479 | -34.7046 | -23.8922
14.754 | -34.7051 | -23.9002
14.7601 | -34.7055 | -23.9085
14.7662 | -34.706 | -23.9167
14.7723 | -34.7064 | -23.9246
14.7784 | -34.7069 | -23.9329
14.7845 | -34.7073 | -23.941
14.7906 | -34.7078 | -23.9493
14.7967 | -34.7082 | -23.9573
14.8028 | -34.7086 | -23.9653
14.8089 | -34.7091 | -23.9734
14.815 | -34.7095 | -23.9813
14.8211 | -34.71 | -23.9893
14.8272 | -34.7104 | -23.9975
14.8334 | -34.7108 | -24.0056
14.8395 | -34.7113 | -24.0136
14.8456 | -34.7118 | -24.0217
14.8517 | -34.7122 | -24.0298
14.8578 | -34.7127 | -24.0377
14.8639 | -34.7131 | -24.0461
14.87 | -34.7136 | -24.0542
14.8761 | -34.7141 | -24.0619
14.8822 | -34.7145 | -24.0701
14.8883 | -34.7149 | -24.0778
14.8944 | -34.7154 | -24.086
14.9005 | -34.7159 | -24.0938
14.9066 | -34.7163 | -24.1019
14.9127 | -34.7168 | -24.1097
14.9188 | -34.7172 | -24.1183
14.9249 | -34.7177 | -24.1264
14.931 | -34.7181 | -24.1344
14.9371 | -34.7186 | -24.1422
14.9432 | -34.719 | -24.1505
14.9493 | -34.7195 | -24.1585
14.9554 | -34.7199 | -24.1666
14.9615 | -34.7204 | -24.1747
14.9676 | -34.7209 | -24.1829
14.9738 | -34.7213 | -24.1908
14.9799 | -34.7218 | -24.1995
14.986 | -34.7223 | -24.2071
14.9921 | -34.7228 | -24.215
14.9982 | -34.7232 | -24.2233
15.0043 | -34.7236 | -24.2313
15.0104 | -34.7241 | -24.2394
15.0165 | -34.7246 | -24.2476
15.0226 | -34.725 | -24.2552
15.0287 | -34.7254 | -24.2637
15.0348 | -34.7259 | -24.2719
15.0409 | -34.7263 | -24.2802
15.047 | -34.7268 | -24.2886
15.0531 | -34.7274 | -24.2968
15.0592 | -34.7278 | -24.3051
15.0653 | -34.7283 | -24.3131
15.0714 | -34.7288 | -24.3208
15.0775 | -34.7293 | -24.3286
15.0836 | -34.7297 | -24.3368
15.0897 | -34.7303 | -24.3447
15.0958 | -34.7307 | -24.3524
15.1019 | -34.7311 | -24.3604
15.108 | -34.7315 | -24.3686
15.1141 | -34.732 | -24.3767
15.1203 | -34.7326 | -24.3847
15.1264 | -34.7331 | -24.3928
15.1325 | -34.7336 | -24.4006
15.1386 | -34.734 | -24.4088
15.1447 | -34.7345 | -24.4167
15.1508 | -34.735 | -24.4251
15.1569 | -34.7355 | -24.4332
15.163 | -34.736 | -24.4408
15.1691 | -34.7365 | -24.4487
15.1752 | -34.737 | -24.4565
15.1813 | -34.7375 | -24.4646
15.1874 | -34.738 | -24.4727
15.1935 | -34.7385 | -24.4803
15.1996 | -34.739 | -24.4885
15.2057 | -34.7394 | -24.4963
15.2118 | -34.7399 | -24.5043
15.2179 | -34.7404 | -24.512
15.224 | -34.7408 | -24.52
15.2301 | -34.7413 | -24.5277
15.2362 | -34.7418 | -24.5357
15.2423 | -34.7423 | -24.544
15.2484 | -34.7427 | -24.5517
15.2545 | -34.7432 | -24.5596
15.2607 | -34.7437 | -24.5675
15.2668 | -34.7443 | -24.5751
15.2729 | -34.7448 | -24.583
15.279 | -34.7453 | -24.5911
15.2851 | -34.7458 | -24.5989
15.2912 | -34.7462 | -24.6067
15.2973 | -34.7467 | -24.6147
15.3034 | -34.7472 | -24.6222
15.3095 | -34.7477 | -24.63
15.3156 | -34.7481 | -24.638
15.3217 | -34.7486 | -24.6457
15.3278 | -34.7491 | -24.6539
15.3339 | -34.7496 | -24.6616
15.34 | -34.7501 | -24.6693
15.3461 | -34.7505 | -24.6773
15.3522 | -34.751 | -24.6851
15.3583 | -34.7515 | -24.6931
15.3644 | -34.752 | -24.7011
15.3705 | -34.7525 | -24.7086
15.3766 | -34.753 | -24.7166
15.3827 | -34.7534 | -24.7242
15.3888 | -34.7539 | -24.7317
15.3949 | -34.7544 | -24.7395
15.401 | -34.7549 | -24.7474
15.4072 | -34.7555 | -24.755
15.4133 | -34.7559 | -24.763
15.4194 | -34.7564 | -24.7708
15.4255 | -34.7569 | -24.7786
15.4316 | -34.7575 | -24.7864
15.4377 | -34.7579 | -24.7942
15.4438 | -34.7584 | -24.8019
15.4499 | -34.7589 | -24.8098
15.456 | -34.7594 | -24.8177
15.4621 | -34.7599 | -24.8257
15.4682 | -34.7604 | -24.8337
15.4743 | -34.761 | -24.8412
15.4804 | -34.7614 | -24.8494
15.4865 | -34.7619 | -24.8573
15.4926 | -34.7624 | -24.8653
15.4987 | -34.7629 | -24.8731
15.5048 | -34.7634 | -24.8808
15.5109 | -34.7639 | -24.8882
15.517 | -34.7644 | -24.8961
15.5231 | -34.765 | -24.9037
15.5292 | -34.7655 | -24.9114
15.5353 | -34.766 | -24.9191
15.5414 | -34.7665 | -24.9269
15.5476 | -34.767 | -24.9347
15.5537 | -34.7675 | -24.9423
15.5598 | -34.768 | -24.9498
15.5659 | -34.7685 | -24.9579
15.572 | -34.7689 | -24.9653
15.5781 | -34.7694 | -24.9734
15.5842 | -34.77 | -24.9808
15.5903 | -34.7705 | -24.9882
15.5964 | -34.771 | -24.9959
15.6025 | -34.7715 | -25.0036
15.6086 | -34.772 | -25.0116
15.6147 | -34.7724 | -25.0194
15.6208 | -34.773 | -25.0269
15.6269 | -34.7736 | -25.0343
15.633 | -34.774 | -25.0419
15.6391 | -34.7745 | -25.0496
15.6452 | -34.775 | -25.0573
15.6513 | -34.7755 | -25.065
15.6574 | -34.776 | -25.0724
15.6635 | -34.7764 | -25.0802
15.6696 | -34.777 | -25.088
15.6757 | -34.7775 | -25.0957
15.6818 | -34.778 | -25.1039
15.688 | -34.7786 | -25.1116
15.6941 | -34.779 | -25.1195
15.7002 | -34.7795 | -25.1271
15.7063 | -34.78 | -25.135
15.7124 | -34.7804 | -25.143
15.7185 | -34.7809 | -25.1507
15.7246 | -34.7814 | -25.1582
15.7307 | -34.782 | -25.166
15.7368 | -34.7825 | -25.1739
15.7429 | -34.783 | -25.1818
15.749 | -34.7835 | -25.1895
15.7551 | -34.784 | -25.1972
15.7612 | -34.7845 | -25.2043
15.7673 | -34.785 | -25.2123
15.7734 | -34.7854 | -25.2199
15.7795 | -34.7859 | -25.2282
15.7856 | -34.7865 | -25.236
15.7917 | -34.787 | -25.2434
15.7978 | -34.7875 | -25.2508
15.8039 | -34.788 | -25.2586
15.81 | -34.7885 | -25.266
15.8161 | -34.789 | -25.2737
15.8222 | -34.7895 | -25.2815
15.8283 | -34.7901 | -25.289
15.8345 | -34.7905 | -25.2966
15.8406 | -34.7911 | -25.3043
15.8467 | -34.7916 | -25.3121
15.8528 | -34.7921 | -25.32
15.8589 | -34.7927 | -25.3274
15.865 | -34.7932 | -25.3349
15.8711 | -34.7937 | -25.3428
15.8772 | -34.7943 | -25.3507
15.8833 | -34.7948 | -25.3588
15.8894 | -34.7953 | -25.3663
15.8955 | -34.7958 | -25.3743
15.9016 | -34.7963 | -25.3818
15.9077 | -34.7968 | -25.3895
15.9138 | -34.7973 | -25.3971
15.9199 | -34.7978 | -25.4048
15.926 | -34.7984 | -25.4126
15.9321 | -34.7989 | -25.4201
15.9382 | -34.7995 | -25.4277
15.9443 | -34.8 | -25.4355
15.9504 | -34.8005 | -25.4433
15.9565 | -34.8011 | -25.451
15.9626 | -34.8016 | -25.4585
15.9687 | -34.8022 | -25.4658
15.9749 | -34.8027 | -25.4737
15.981 | -34.8033 | -25.4812
15.9871 | -34.8038 | -25.4886
15.9932 | -34.8043 | -25.4964
15.9993 | -34.8048 | -25.5036
16.0054 | -34.8054 | -25.5109
16.0115 | -34.8059 | -25.5185
16.0176 | -34.8064 | -25.5263
16.0237 | -34.8069 | -25.5344
16.0298 | -34.8074 | -25.5414
16.0359 | -34.808 | -25.5488
16.042 | -34.8085 | -25.5564
16.0481 | -34.809 | -25.5641
16.0542 | -34.8096 | -25.5718
16.0603 | -34.8101 | -25.5793
16.0664 | -34.8106 | -25.587
16.0725 | -34.8112 | -25.5942
16.0786 | -34.8118 | -25.6024
16.0847 | -34.8122 | -25.61
16.0908 | -34.8128 | -25.6178
16.0969 | -34.8133 | -25.6253
16.103 | -34.8138 | -25.6331
16.1091 | -34.8143 | -25.6407
16.1152 | -34.8148 | -25.6485
16.1214 | -34.8154 | -25.656
16.1275 | -34.816 | -25.6634
16.1336 | -34.8165 | -25.6711
16.1397 | -34.8171 | -25.6788
16.1458 | -34.8176 | -25.6862
16.1519 | -34.8181 | -25.6939
16.158 | -34.8187 | -25.7016
16.1641 | -34.8192 | -25.7095
16.1702 | -34.8197 | -25.7171
16.1763 | -34.8202 | -25.7244
16.1824 | -34.8207 | -25.7316
16.1885 | -34.8213 | -25.7392
16.1946 | -34.8218 | -25.7468
16.2007 | -34.8223 | -25.7548
16.2068 | -34.8228 | -25.7623
16.2129 | -34.8233 | -25.7698
16.219 | -34.8238 | -25.7771
16.2251 | -34.8244 | -25.7849
16.2312 | -34.8248 | -25.7925
16.2373 | -34.8254 | -25.8
16.2434 | -34.8259 | -25.8074
16.2495 | -34.8264 | -25.8149
16.2556 | -34.8271 | -25.8226
16.2618 | -34.8276 | -25.83
16.2679 | -34.8281 | -25.8377
16.274 | -34.8286 | -25.8452
16.2801 | -34.8291 | -25.8527
16.2862 | -34.8297 | -25.8603
16.2923 | -34.8302 | -25.8681
16.2984 | -34.8307 | -25.8758
16.3045 | -34.8313 | -25.8834
16.3106 | -34.8318 | -25.8907
16.3167 | -34.8323 | -25.8986
16.3228 | -34.8329 | -25.9065
16.3289 | -34.8334 | -25.9141
16.335 | -34.8339 | -25.9218
16.3411 | -34.8343 | -25.9289
16.3472 | -34.8349 | -25.9366
16.3533 | -34.8354 | -25.9444
16.3594 | -34.8359 | -25.9516
16.3655 | -34.8364 | -25.9593
16.3716 | -34.837 | -25.9669
16.3777 | -34.8375 | -25.9744
16.3838 | -34.838 | -25.9819
16.3899 | -34.8385 | -25.9894
16.396 | -34.839 | -25.9969
16.4021 | -34.8395 | -26.0043
16.4083 | -34.8401 | -26.0118
16.4144 | -34.8406 | -26.0193
16.4205 | -34.8412 | -26.0269
16.4266 | -34.8417 | -26.0342
16.4327 | -34.8423 | -26.0416
16.4388 | -34.8428 | -26.0492
16.4449 | -34.8433 | -26.0567
16.451 | -34.8438 | -26.0644
16.4571 | -34.8443 | -26.0721
16.4632 | -34.8448 | -26.0795
16.4693 | -34.8453 | -26.0871
16.4754 | -34.8459 | -26.0946
16.4815 | -34.8464 | -26.1024
16.4876 | -34.847 | -26.11
16.4937 | -34.8475 | -26.1173
16.4998 | -34.848 | -26.1249
16.5059 | -34.8486 | -26.1325
16.512 | -34.8491 | -26.1403
16.5181 | -34.8497 | -26.1479
16.5242 | -34.8502 | -26.1555
16.5303 | -34.8507 | -26.1629
16.5364 | -34.8512 | -26.1702
16.5425 | -34.8517 | -26.1779
16.5487 | -34.8523 | -26.1853
16.5548 | -34.8528 | -26.1927
16.5609 | -34.8533 | -26.2003
16.567 | -34.8538 | -26.2079
16.5731 | -34.8544 | -26.2154
16.5792 | -34.8549 | -26.2229
16.5853 | -34.8555 | -26.2303
16.5914 | -34.856 | -26.2378
16.5975 | -34.8565 | -26.2453
16.6036 | -34.857 | -26.2529
16.6097 | -34.8576 | -26.2604
16.6158 | -34.858 | -26.2681
16.6219 | -34.8586 | -26.2756
16.628 | -34.859 | -26.2832
16.6341 | -34.8596 | -26.2905
16.6402 | -34.8602 | -26.2974
16.6463 | -34.8607 | -26.3049
16.6524 | -34.8612 | -26.3124
16.6585 | -34.8617 | -26.32
16.6646 | -34.8623 | -26.3274
16.6707 | -34.8628 | -26.3348
16.6768 | -34.8633 | -26.3423
16.6829 | -34.8638 | -26.3498
16.689 | -34.8643 | -26.3575
16.6952 | -34.8648 | -26.365
16.7013 | -34.8654 | -26.3726
16.7074 | -34.8659 | -26.3801
16.7135 | -34.8664 | -26.3874
16.7196 | -34.867 | -26.395
16.7257 | -34.8675 | -26.4023
16.7318 | -34.868 | -26.4099
16.7379 | -34.8685 | -26.4177
16.744 | -34.8691 | -26.4251
16.7501 | -34.8696 | -26.4327
16.7562 | -34.8702 | -26.4402
16.7623 | -34.8707 | -26.4475
16.7684 | -34.8712 | -26.4552
16.7745 | -34.8717 | -26.4626
16.7806 | -34.8722 | -26.4702
16.7867 | -34.8727 | -26.4777
16.7928 | -34.8732 | -26.4853
16.7989 | -34.8737 | -26.4926
16.805 | -34.8742 | -26.5001
16.8111 | -34.8748 | -26.5079
16.8172 | -34.8753 | -26.5157
16.8233 | -34.8759 | -26.5234
16.8294 | -34.8764 | -26.5308
16.8356 | -34.8769 | -26.5385
16.8417 | -34.8774 | -26.5459
16.8478 | -34.8779 | -26.5533
16.8539 | -34.8785 | -26.5605
16.86 | -34.879 | -26.5679
16.8661 | -34.8795 | -26.5755
16.8722 | -34.8801 | -26.583
16.8783 | -34.8806 | -26.5902
16.8844 | -34.8811 | -26.5976
16.8905 | -34.8817 | -26.6053
16.8966 | -34.8822 | -26.613
16.9027 | -34.8827 | -26.6202
16.9088 | -34.8832 | -26.6281
16.9149 | -34.8838 | -26.6358
16.921 | -34.8843 | -26.6433
16.9271 | -34.8848 | -26.6512
16.9332 | -34.8854 | -26.6589
16.9393 | -34.8859 | -26.6667
16.9454 | -34.8863 | -26.6739
16.9515 | -34.8868 | -26.6817
16.9576 | -34.8873 | -26.6892
16.9637 | -34.8879 | -26.6965
16.9698 | -34.8885 | -26.7037
16.9759 | -34.889 | -26.711
16.9821 | -34.8895 | -26.7187
16.9882 | -34.89 | -26.7264
16.9943 | -34.8906 | -26.734
17.0004 | -34.8911 | -26.7415
17.0065 | -34.8917 | -26.7491
17.0126 | -34.8922 | -26.7566
17.0187 | -34.8927 | -26.7642
17.0248 | -34.8932 | -26.7716
17.0309 | -34.8938 | -26.779
17.037 | -34.8943 | -26.7865
17.0431 | -34.8948 | -26.7938
17.0492 | -34.8953 | -26.8012
17.0553 | -34.8958 | -26.8086
17.0614 | -34.8963 | -26.816
17.0675 | -34.8968 | -26.8236
17.0736 | -34.8973 | -26.8312
17.0797 | -34.8978 | -26.8383
17.0858 | -34.8984 | -26.8459
17.0919 | -34.8989 | -26.8535
17.098 | -34.8994 | -26.861
17.1041 | -34.8999 | -26.8682
17.1102 | -34.9005 | -26.8761
17.1163 | -34.901 | -26.8837
17.1225 | -34.9016 | -26.8909
17.1286 | -34.9021 | -26.8983
17.1347 | -34.9025 | -26.9058
17.1408 | -34.9031 | -26.9134
17.1469 | -34.9036 | -26.9208
17.153 | -34.9041 | -26.9282
17.1591 | -34.9046 | -26.9354
17.1652 | -34.9051 | -26.9427
17.1713 | -34.9056 | -26.9501
17.1774 | -34.9061 | -26.9577
17.1835 | -34.9067 | -26.9653
17.1896 | -34.9072 | -26.9724
17.1957 | -34.9077 | -26.9799
17.2018 | -34.9081 | -26.9876
17.2079 | -34.9087 | -26.995
17.214 | -34.9092 | -27.0025
17.2201 | -34.9097 | -27.0099
17.2262 | -34.9101 | -27.0176
17.2323 | -34.9107 | -27.0252
17.2384 | -34.9112 | -27.0326
17.2445 | -34.9117 | -27.0399
17.2506 | -34.9122 | -27.0475
17.2567 | -34.9127 | -27.0549
17.2628 | -34.9132 | -27.0627
17.269 | -34.9137 | -27.0702
17.2751 | -34.9143 | -27.0777
17.2812 | -34.9148 | -27.0851
17.2873 | -34.9153 | -27.0927
17.2934 | -34.9158 | -27.1
17.2995 | -34.9163 | -27.1074
17.3056 | -34.9167 | -27.1147
17.3117 | -34.9173 | -27.1222
17.3178 | -34.9178 | -27.1296
17.3239 | -34.9182 | -27.137
17.33 | -34.9187 | -27.1444
17.3361 | -34.9192 | -27.1518
17.3422 | -34.9197 | -27.1591
17.3483 | -34.9202 | -27.1668
17.3544 | -34.9207 | -27.1742
17.3605 | -34.9212 | -27.1817
17.3666 | -34.9218 | -27.1893
17.3727 | -34.9223 | -27.1967
17.3788 | -34.9228 | -27.204
17.3849 | -34.9233 | -27.2114
17.391 | -34.9238 | -27.2189
17.3971 | -34.9243 | -27.2263
17.4032 | -34.9248 | -27.234
17.4094 | -34.9253 | -27.2414
17.4155 | -34.9258 | -27.2486
17.4216 | -34.9263 | -27.2561
17.4277 | -34.9268 | -27.2637
17.4338 | -34.9273 | -27.2712
17.4399 | -34.9278 | -27.2786
17.446 | -34.9283 | -27.2861
17.4521 | -34.9288 | -27.2934
17.4582 | -34.9293 | -27.3007
17.4643 | -34.9298 | -27.3083
17.4704 | -34.9303 | -27.3157
17.4765 | -34.9308 | -27.3232
17.4826 | -34.9313 | -27.3305
17.4887 | -34.9318 | -27.3383
17.4948 | -34.9323 | -27.3457
17.5009 | -34.9328 | -27.353
17.507 | -34.9333 | -27.3606
17.5131 | -34.9338 | -27.368
17.5192 | -34.9343 | -27.3753
17.5253 | -34.9347 | -27.3829
17.5314 | -34.9352 | -27.3903
17.5375 | -34.9357 | -27.3977
17.5436 | -34.9361 | -27.4052
17.5497 | -34.9366 | -27.4128
17.5559 | -34.9371 | -27.4201
17.562 | -34.9376 | -27.4273
17.5681 | -34.9382 | -27.4346
17.5742 | -34.9386 | -27.4422
17.5803 | -34.9392 | -27.4495
17.5864 | -34.9397 | -27.4567
17.5925 | -34.9402 | -27.4646
17.5986 | -34.9408 | -27.4717
17.6047 | -34.9413 | -27.4795
17.6108 | -34.9418 | -27.4867
17.6169 | -34.9423 | -27.4941
17.623 | -34.9428 | -27.5013
17.6291 | -34.9433 | -27.5087
17.6352 | -34.9438 | -27.5158
17.6413 | -34.9443 | -27.5234
17.6474 | -34.9448 | -27.531
17.6535 | -34.9453 | -27.5383
17.6596 | -34.9458 | -27.5456
17.6657 | -34.9463 | -27.553
17.6718 | -34.9468 | -27.5605
17.6779 | -34.9473 | -27.5678
17.684 | -34.9478 | -27.5754
17.6901 | -34.9483 | -27.5828
17.6963 | -34.9488 | -27.5901
17.7024 | -34.9493 | -27.5974
17.7085 | -34.9497 | -27.6047
17.7146 | -34.9503 | -27.6121
17.7207 | -34.9507 | -27.6193
17.7268 | -34.9513 | -27.6267
17.7329 | -34.9518 | -27.6343
17.739 | -34.9523 | -27.6418
17.7451 | -34.9528 | -27.649
17.7512 | -34.9533 | -27.6565
17.7573 | -34.9537 | -27.6639
17.7634 | -34.9542 | -27.6711
17.7695 | -34.9547 | -27.6786
17.7756 | -34.9552 | -27.6862
17.7817 | -34.9557 | -27.6936
17.7878 | -34.9562 | -27.701
17.7939 | -34.9567 | -27.7083
17.8 | -34.9572 | -27.7156
17.8061 | -34.9577 | -27.723
17.8122 | -34.9582 | -27.7305
17.8183 | -34.9586 | -27.7379
17.8244 | -34.9592 | -27.7453
17.8305 | -34.9597 | -27.7527
17.8366 | -34.9602 | -27.7596
17.8428 | -34.9606 | -27.7673
17.8489 | -34.9611 | -27.7745
17.855 | -34.9616 | -27.7817
17.8611 | -34.9621 | -27.7888
17.8672 | -34.9626 | -27.796
17.8733 | -34.9631 | -27.8037
17.8794 | -34.9636 | -27.8113
17.8855 | -34.9641 | -27.8187
17.8916 | -34.9646 | -27.826
17.8977 | -34.9651 | -27.8333
17.9038 | -34.9656 | -27.8404
17.9099 | -34.9661 | -27.8478
17.916 | -34.9666 | -27.8552
17.9221 | -34.967 | -27.8626
17.9282 | -34.9675 | -27.87
17.9343 | -34.968 | -27.8776
17.9404 | -34.9685 | -27.8848
17.9465 | -34.969 | -27.8923
17.9526 | -34.9695 | -27.9
17.9587 | -34.97 | -27.9073
17.9648 | -34.9705 | -27.9146
17.9709 | -34.971 | -27.922
17.977 | -34.9715 | -27.9293
17.9832 | -34.9719 | -27.9367
17.9893 | -34.9724 | -27.9443
17.9954 | -34.9729 | -27.9519
18.0015 | -34.9735 | -27.9595
18.0076 | -34.9739 | -27.9663
18.0137 | -34.9744 | -27.974
18.0198 | -34.9748 | -27.9812
18.0259 | -34.9753 | -27.9887
18.032 | -34.9758 | -27.9959
18.0381 | -34.9763 | -28.0031
18.0442 | -34.9768 | -28.0105
18.0503 | -34.9773 | -28.018
18.0564 | -34.9777 | -28.0254
18.0625 | -34.9782 | -28.0327
18.0686 | -34.9788 | -28.04
18.0747 | -34.9792 | -28.0473
18.0808 | -34.9797 | -28.0545
18.0869 | -34.9803 | -28.0621
18.093 | -34.9808 | -28.0694
18.0991 | -34.9812 | -28.077
18.1052 | -34.9817 | -28.0845
18.1113 | -34.9822 | -28.0915
18.1174 | -34.9827 | -28.0987
18.1236 | -34.9832 | -28.1062
18.1297 | -34.9837 | -28.1134
18.1358 | -34.9843 | -28.1208
18.1419 | -34.9848 | -28.1283
18.148 | -34.9853 | -28.1357
18.1541 | -34.9859 | -28.1431
18.1602 | -34.9864 | -28.1506
18.1663 | -34.9869 | -28.1579
18.1724 | -34.9874 | -28.1655
18.1785 | -34.9879 | -28.1731
18.1846 | -34.9884 | -28.1802
18.1907 | -34.9889 | -28.1879
18.1968 | -34.9894 | -28.1954
18.2029 | -34.9898 | -28.2022
18.209 | -34.9904 | -28.2099
18.2151 | -34.9909 | -28.2175
18.2212 | -34.9914 | -28.2246
18.2273 | -34.9919 | -28.2321
18.2334 | -34.9924 | -28.2393
18.2395 | -34.9928 | -28.2465
18.2456 | -34.9934 | -28.2538
18.2517 | -34.9938 | -28.2612
18.2578 | -34.9943 | -28.2685
18.2639 | -34.9948 | -28.2761
18.2701 | -34.9954 | -28.2834
18.2762 | -34.9959 | -28.2909
18.2823 | -34.9964 | -28.2982
18.2884 | -34.9969 | -28.3058
18.2945 | -34.9973 | -28.3128
18.3006 | -34.9978 | -28.3202
18.3067 | -34.9983 | -28.3274
18.3128 | -34.9988 | -28.3347
18.3189 | -34.9993 | -28.3422
18.325 | -34.9998 | -28.349
18.3311 | -35.0003 | -28.3564
18.3372 | -35.0008 | -28.3636
18.3433 | -35.0012 | -28.3707
18.3494 | -35.0017 | -28.3781
18.3555 | -35.0022 | -28.3852
18.3616 | -35.0027 | -28.3924
18.3677 | -35.0032 | -28.3995
18.3738 | -35.0037 | -28.4066
18.3799 | -35.0042 | -28.4137
18.386 | -35.0046 | -28.4208
18.3921 | -35.0051 | -28.4281
18.3982 | -35.0056 | -28.4353
18.4043 | -35.006 | -28.4424
18.4105 | -35.0065 | -28.4498
18.4166 | -35.007 | -28.4569
18.4227 | -35.0075 | -28.4642
18.4288 | -35.008 | -28.4716
18.4349 | -35.0085 | -28.4791
18.441 | -35.009 | -28.4861
18.4471 | -35.0095 | -28.4935
18.4532 | -35.01 | -28.5008
18.4593 | -35.0104 | -28.508
18.4654 | -35.0109 | -28.5151
18.4715 | -35.0114 | -28.5224
18.4776 | -35.0119 | -28.5299
18.4837 | -35.0124 | -28.5373
18.4898 | -35.0129 | -28.545
18.4959 | -35.0135 | -28.5523
18.502 | -35.014 | -28.5596
18.5081 | -35.0144 | -28.5669
18.5142 | -35.0148 | -28.5744
18.5203 | -35.0154 | -28.5814
18.5264 | -35.0158 | -28.5884
18.5325 | -35.0163 | -28.5958
18.5386 | -35.0169 | -28.6034
18.5447 | -35.0174 | -28.6103
18.5508 | -35.0179 | -28.6174
18.557 | -35.0184 | -28.6248
18.5631 | -35.0189 | -28.6321
18.5692 | -35.0194 | -28.6394
18.5753 | -35.0198 | -28.6465
18.5814 | -35.0203 | -28.654
18.5875 | -35.0208 | -28.6617
18.5936 | -35.0213 | -28.6693
18.5997 | -35.0217 | -28.6767
18.6058 | -35.0222 | -28.6842
18.6119 | -35.0227 | -28.6918
18.618 | -35.0232 | -28.6991
18.6241 | -35.0237 | -28.7066
18.6302 | -35.0242 | -28.7141
18.6363 | -35.0247 | -28.7213
18.6424 | -35.0252 | -28.7285
18.6485 | -35.0256 | -28.7359
18.6546 | -35.0262 | -28.743
18.6607 | -35.0267 | -28.7502
18.6668 | -35.0272 | -28.7574
18.6729 | -35.0277 | -28.7647
18.679 | -35.0282 | -28.7717
18.6851 | -35.0287 | -28.7788
18.6912 | -35.0292 | -28.786
18.6974 | -35.0297 | -28.7931
18.7035 | -35.0301 | -28.8003
18.7096 | -35.0306 | -28.8077
18.7157 | -35.0312 | -28.8147
18.7218 | -35.0316 | -28.8222
18.7279 | -35.0322 | -28.8295
18.734 | -35.0327 | -28.8367
18.7401 | -35.0332 | -28.8436
18.7462 | -35.0337 | -28.8503
18.7523 | -35.0342 | -28.8573
18.7584 | -35.0347 | -28.8647
18.7645 | -35.0352 | -28.8718
18.7706 | -35.0357 | -28.8792
18.7767 | -35.0362 | -28.8865
18.7828 | -35.0367 | -28.8933
18.7889 | -35.0373 | -28.9005
18.795 | -35.0378 | -28.9077
18.8011 | -35.0383 | -28.9152
18.8072 | -35.0388 | -28.9225
18.8133 | -35.0393 | -28.9298
18.8194 | -35.0397 | -28.9368
18.8255 | -35.0403 | -28.9441
18.8316 | -35.0407 | -28.9513
18.8377 | -35.0412 | -28.9584
18.8439 | -35.0417 | -28.9657
18.85 | -35.0422 | -28.9729
18.8561 | -35.0427 | -28.9802
18.8622 | -35.0432 | -28.9877
18.8683 | -35.0437 | -28.9947
18.8744 | -35.0442 | -29.0019
18.8805 | -35.0447 | -29.0091
18.8866 | -35.0452 | -29.0166
18.8927 | -35.0457 | -29.0237
18.8988 | -35.0461 | -29.0311
18.9049 | -35.0466 | -29.0382
18.911 | -35.0472 | -29.0454
18.9171 | -35.0477 | -29.0526
18.9232 | -35.0482 | -29.0596
18.9293 | -35.0486 | -29.0668
18.9354 | -35.0491 | -29.0739
18.9415 | -35.0496 | -29.0808
18.9476 | -35.0501 | -29.0882
18.9537 | -35.0506 | -29.0956
18.9598 | -35.0511 | -29.1022
18.9659 | -35.0517 | -29.1094
18.972 | -35.0522 | -29.1167
18.9781 | -35.0527 | -29.1238
18.9843 | -35.0532 | -29.1311
18.9904 | -35.0536 | -29.1384
18.9965 | -35.0542 | -29.1454
19.0026 | -35.0547 | -29.1527
19.0087 | -35.0552 | -29.1599
19.0148 | -35.0558 | -29.1668
19.0209 | -35.0562 | -29.1745
19.027 | -35.0568 | -29.1817
19.0331 | -35.0573 | -29.1889
19.0392 | -35.0578 | -29.1963
19.0453 | -35.0583 | -29.2035
19.0514 | -35.0588 | -29.2106
19.0575 | -35.0593 | -29.2178
19.0636 | -35.0598 | -29.2243
19.0697 | -35.0603 | -29.2317
19.0758 | -35.0608 | -29.2385
19.0819 | -35.0613 | -29.2461
19.088 | -35.0618 | -29.2528
19.0941 | -35.0623 | -29.2598
19.1002 | -35.0628 | -29.2666
19.1063 | -35.0634 | -29.2738
19.1124 | -35.0639 | -29.281
19.1185 | -35.0644 | -29.2882
19.1246 | -35.0649 | -29.2952
19.1308 | -35.0654 | -29.302
19.1369 | -35.066 | -29.309
19.143 | -35.0664 | -29.3162
19.1491 | -35.0669 | -29.3232
19.1552 | -35.0674 | -29.3305
19.1613 | -35.068 | -29.3374
19.1674 | -35.0685 | -29.3448
19.1735 | -35.0689 | -29.3518
19.1796 | -35.0695 | -29.3589
19.1857 | -35.0699 | -29.3661
19.1918 | -35.0704 | -29.3732
19.1979 | -35.0709 | -29.3804
19.204 | -35.0715 | -29.3873
19.2101 | -35.072 | -29.3946
19.2162 | -35.0724 | -29.4019
19.2223 | -35.0729 | -29.4085
19.2284 | -35.0735 | -29.4159
19.2345 | -35.074 | -29.4232
19.2406 | -35.0745 | -29.4302
19.2467 | -35.075 | -29.4375
19.2528 | -35.0755 | -29.4446
19.2589 | -35.076 | -29.4519
19.265 | -35.0765 | -29.459
19.2712 | -35.077 | -29.4662
19.2773 | -35.0775 | -29.473
19.2834 | -35.078 | -29.4798
19.2895 | -35.0785 | -29.4869
19.2956 | -35.079 | -29.4945
19.3017 | -35.0795 | -29.5013
19.3078 | -35.08 | -29.5087
19.3139 | -35.0805 | -29.5158
19.32 | -35.0811 | -29.5229
19.3261 | -35.0815 | -29.53
19.3322 | -35.082 | -29.537
19.3383 | -35.0826 | -29.544
19.3444 | -35.0831 | -29.5512
19.3505 | -35.0836 | -29.5583
19.3566 | -35.0841 | -29.5654
19.3627 | -35.0846 | -29.5726
19.3688 | -35.0851 | -29.5799
19.3749 | -35.0857 | -29.5872
19.381 | -35.0862 | -29.5938
19.3871 | -35.0867 | -29.6007
19.3932 | -35.0872 | -29.6079
19.3993 | -35.0876 | -29.6152
19.4054 | -35.0882 | -29.6227
19.4115 | -35.0887 | -29.6298
19.4177 | -35.0892 | -29.6368
19.4238 | -35.0897 | -29.6437
19.4299 | -35.0902 | -29.6506
19.436 | -35.0907 | -29.6583
19.4421 | -35.0913 | -29.6648
19.4482 | -35.0918 | -29.6719
19.4543 | -35.0923 | -29.6787
19.4604 | -35.0928 | -29.6854
19.4665 | -35.0934 | -29.6924
19.4726 | -35.0939 | -29.6992
19.4787 | -35.0944 | -29.7062
19.4848 | -35.0948 | -29.7134
19.4909 | -35.0954 | -29.7205
19.497 | -35.0959 | -29.7274
19.5031 | -35.0964 | -29.7343
19.5092 | -35.0968 | -29.7415
19.5153 | -35.0973 | -29.7485
19.5214 | -35.0979 | -29.7553
19.5275 | -35.0983 | -29.7624
19.5336 | -35.0988 | -29.7694
19.5397 | -35.0993 | -29.7768
19.5458 | -35.0998 | -29.7838
19.5519 | -35.1003 | -29.7906
19.5581 | -35.1009 | -29.7976
19.5642 | -35.1014 | -29.8045
19.5703 | -35.102 | -29.8115
19.5764 | -35.1025 | -29.8187
19.5825 | -35.1029 | -29.826
19.5886 | -35.1034 | -29.8328
19.5947 | -35.1039 | -29.8397
19.6008 | -35.1044 | -29.8466
19.6069 | -35.1049 | -29.8536
19.613 | -35.1054 | -29.8611
19.6191 | -35.1058 | -29.8682
19.6252 | -35.1064 | -29.8753
19.6313 | -35.1069 | -29.8823
19.6374 | -35.1074 | -29.8893
19.6435 | -35.1079 | -29.8962
19.6496 | -35.1084 | -29.9031
19.6557 | -35.109 | -29.9101
19.6618 | -35.1095 | -29.9168
19.6679 | -35.11 | -29.9238
19.674 | -35.1105 | -29.9307
19.6801 | -35.111 | -29.9376
19.6862 | -35.1116 | -29.9448
19.6923 | -35.1121 | -29.9516
19.6984 | -35.1126 | -29.9585
19.7046 | -35.1132 | -29.9654
19.7107 | -35.1137 | -29.9722
19.7168 | -35.1141 | -29.9795
19.7229 | -35.1146 | -29.9867
19.729 | -35.1151 | -29.9941
19.7351 | -35.1156 | -30.0006
19.7412 | -35.1161 | -30.0076
19.7473 | -35.1166 | -30.0146
19.7534 | -35.1171 | -30.0215
19.7595 | -35.1176 | -30.0287
19.7656 | -35.1181 | -30.036
19.7717 | -35.1186 | -30.0428
19.7778 | -35.1191 | -30.05
19.7839 | -35.1196 | -30.0568
19.79 | -35.1201 | -30.0638
19.7961 | -35.1206 | -30.0709
19.8022 | -35.1211 | -30.0776
19.8083 | -35.1216 | -30.0846
19.8144 | -35.1221 | -30.0915
19.8205 | -35.1226 | -30.0986
19.8266 | -35.1232 | -30.1055
19.8327 | -35.1237 | -30.1125
19.8388 | -35.1242 | -30.1195
19.845 | -35.1247 | -30.1262
19.8511 | -35.1252 | -30.1336
19.8572 | -35.1257 | -30.1406
19.8633 | -35.1262 | -30.1474
19.8694 | -35.1268 | -30.1546
19.8755 | -35.1273 | -30.1614
19.8816 | -35.1278 | -30.1687
19.8877 | -35.1283 | -30.1756
19.8938 | -35.1288 | -30.1825
19.8999 | -35.1294 | -30.1892
19.906 | -35.1299 | -30.1964
19.9121 | -35.1304 | -30.2031
19.9182 | -35.131 | -30.2098
19.9243 | -35.1314 | -30.2167
19.9304 | -35.1319 | -30.2239
19.9365 | -35.1324 | -30.2311
19.9426 | -35.1329 | -30.2378
19.9487 | -35.1334 | -30.2444
19.9548 | -35.1339 | -30.2515
19.9609 | -35.1345 | -30.2583
19.967 | -35.135 | -30.2651
19.9731 | -35.1355 | -30.2721
19.9792 | -35.1359 | -30.2786
19.9853 | -35.1364 | -30.2855
19.9915 | -35.137 | -30.2925
19.9976 | -35.1375 | -30.2994
20.0037 | -35.138 | -30.3065
20.0098 | -35.1385 | -30.3134
20.0159 | -35.1391 | -30.3204
20.022 | -35.1396 | -30.3277
20.0281 | -35.1401 | -30.3347
20.0342 | -35.1406 | -30.3415
20.0403 | -35.1411 | -30.3481
20.0464 | -35.1416 | -30.3555
20.0525 | -35.1421 | -30.3618
20.0586 | -35.1426 | -30.3689
20.0647 | -35.1431 | -30.3764
20.0708 | -35.1437 | -30.3835
20.0769 | -35.1442 | -30.3908
20.083 | -35.1447 | -30.3976
20.0891 | -35.1452 | -30.4045
20.0952 | -35.1457 | -30.4111
20.1013 | -35.1462 | -30.4178
20.1074 | -35.1467 | -30.4251
20.1135 | -35.1473 | -30.4319
20.1196 | -35.1477 | -30.4388
20.1257 | -35.1482 | -30.4459
20.1319 | -35.1488 | -30.453
20.138 | -35.1493 | -30.4596
20.1441 | -35.1498 | -30.4664
20.1502 | -35.1503 | -30.4738
20.1563 | -35.1508 | -30.4803
20.1624 | -35.1513 | -30.4871
20.1685 | -35.1518 | -30.4942
20.1746 | -35.1524 | -30.5013
20.1807 | -35.1529 | -30.5081
20.1868 | -35.1533 | -30.5151
20.1929 | -35.1539 | -30.5216
20.199 | -35.1543 | -30.5289
20.2051 | -35.1548 | -30.5357
20.2112 | -35.1553 | -30.5426
20.2173 | -35.1558 | -30.5497
20.2234 | -35.1563 | -30.5568
20.2295 | -35.1568 | -30.5632
20.2356 | -35.1573 | -30.5701
20.2417 | -35.1578 | -30.5767
20.2478 | -35.1583 | -30.5835
20.2539 | -35.1588 | -30.5907
20.26 | -35.1593 | -30.5978
20.2661 | -35.1598 | -30.6047
20.2723 | -35.1602 | -30.6117
20.2784 | -35.1608 | -30.6186
20.2845 | -35.1613 | -30.6255
20.2906 | -35.1619 | -30.6319
20.2967 | -35.1624 | -30.6387
20.3028 | -35.1628 | -30.6456
20.3089 | -35.1633 | -30.6527
20.315 | -35.1638 | -30.6596
20.3211 | -35.1643 | -30.6665
20.3272 | -35.1648 | -30.6734
20.3333 | -35.1653 | -30.6804
20.3394 | -35.1658 | -30.6871
20.3455 | -35.1663 | -30.694
20.3516 | -35.1668 | -30.701
20.3577 | -35.1673 | -30.708
20.3638 | -35.1678 | -30.7152
20.3699 | -35.1684 | -30.7216
20.376 | -35.1689 | -30.7286
20.3821 | -35.1694 | -30.7353
20.3882 | -35.1699 | -30.7422
20.3943 | -35.1705 | -30.7491
20.4004 | -35.171 | -30.7555
20.4065 | -35.1715 | -30.7624
20.4126 | -35.172 | -30.7692
20.4188 | -35.1725 | -30.776
20.4249 | -35.173 | -30.7828
20.431 | -35.1735 | -30.7899
20.4371 | -35.174 | -30.7963
20.4432 | -35.1745 | -30.8031
20.4493 | -35.175 | -30.8102
20.4554 | -35.1756 | -30.817
20.4615 | -35.1761 | -30.8241
20.4676 | -35.1766 | -30.8312
20.4737 | -35.1771 | -30.8377
20.4798 | -35.1777 | -30.8445
20.4859 | -35.1782 | -30.8511
20.492 | -35.1787 | -30.858
20.4981 | -35.1793 | -30.8648
20.5042 | -35.1798 | -30.8715
20.5103 | -35.1802 | -30.8778
20.5164 | -35.1808 | -30.8848
20.5225 | -35.1812 | -30.8921
20.5286 | -35.1817 | -30.8992
20.5347 | -35.1821 | -30.9059
20.5408 | -35.1827 | -30.9127
20.5469 | -35.1832 | -30.9192
20.553 | -35.1837 | -30.9264
20.5592 | -35.1842 | -30.9334
20.5653 | -35.1846 | -30.9404
20.5714 | -35.1852 | -30.9468
20.5775 | -35.1857 | -30.9531
20.5836 | -35.1862 | -30.96
20.5897 | -35.1868 | -30.9666
20.5958 | -35.1873 | -30.9735
20.6019 | -35.1878 | -30.9806
20.608 | -35.1883 | -30.9873
20.6141 | -35.1888 | -30.9938
20.6202 | -35.1893 | -31.0007
20.6263 | -35.1898 | -31.0075
20.6324 | -35.1903 | -31.0139
20.6385 | -35.1909 | -31.0208
20.6446 | -35.1913 | -31.028
20.6507 | -35.1918 | -31.0347
20.6568 | -35.1923 | -31.0414
20.6629 | -35.1928 | -31.0484
20.669 | -35.1933 | -31.0549
20.6751 | -35.1938 | -31.0616
20.6812 | -35.1943 | -31.0678
20.6873 | -35.1948 | -31.0746
20.6934 | -35.1953 | -31.0816
20.6995 | -35.1958 | -31.0881
20.7057 | -35.1964 | -31.0949
20.7118 | -35.1969 | -31.1016
20.7179 | -35.1974 | -31.1081
20.724 | -35.1979 | -31.1147
20.7301 | -35.1984 | -31.1216
20.7362 | -35.1989 | -31.1287
20.7423 | -35.1994 | -31.1353
20.7484 | -35.1999 | -31.142
20.7545 | -35.2005 | -31.1485
20.7606 | -35.201 | -31.1552
20.7667 | -35.2015 | -31.162
20.7728 | -35.202 | -31.1687
20.7789 | -35.2025 | -31.1755
20.785 | -35.203 | -31.1822
20.7911 | -35.2035 | -31.189
20.7972 | -35.2039 | -31.1959
20.8033 | -35.2044 | -31.2021
20.8094 | -35.205 | -31.2093
20.8155 | -35.2054 | -31.216
20.8216 | -35.2059 | -31.2229
20.8277 | -35.2065 | -31.2296
20.8338 | -35.207 | -31.2365
20.8399 | -35.2076 | -31.2431
20.8461 | -35.2081 | -31.2499
20.8522 | -35.2087 | -31.2568
20.8583 | -35.2092 | -31.2635
20.8644 | -35.2097 | -31.27
20.8705 | -35.2103 | -31.2773
20.8766 | -35.2107 | -31.2843
20.8827 | -35.2113 | -31.2908
20.8888 | -35.2118 | -31.2973
20.8949 | -35.2123 | -31.3037
20.901 | -35.2128 | -31.3103
20.9071 | -35.2133 | -31.317
20.9132 | -35.2138 | -31.3238
20.9193 | -35.2143 | -31.3305
20.9254 | -35.2148 | -31.3373
20.9315 | -35.2153 | -31.3438
20.9376 | -35.2158 | -31.3509
20.9437 | -35.2163 | -31.3578
20.9498 | -35.2168 | -31.3642
20.9559 | -35.2174 | -31.3705
20.962 | -35.2179 | -31.3776
20.9681 | -35.2183 | -31.384
20.9742 | -35.2188 | -31.3906
20.9803 | -35.2193 | -31.397
20.9864 | -35.2198 | -31.4035
20.9926 | -35.2203 | -31.4107
20.9987 | -35.2208 | -31.4169
21.0048 | -35.2212 | -31.4241
21.0109 | -35.2217 | -31.4309
21.017 | -35.2222 | -31.4373
21.0231 | -35.2227 | -31.4442
21.0292 | -35.2232 | -31.4511
21.0353 | -35.2236 | -31.4578
21.0414 | -35.2241 | -31.4642
21.0475 | -35.2246 | -31.4708
21.0536 | -35.2251 | -31.4771
21.0597 | -35.2256 | -31.4838
21.0658 | -35.2261 | -31.4901
21.0719 | -35.2265 | -31.4966
21.078 | -35.2271 | -31.5034
21.0841 | -35.2275 | -31.5097
21.0902 | -35.228 | -31.5161
21.0963 | -35.2285 | -31.5228
21.1024 | -35.229 | -31.5303
21.1085 | -35.2294 | -31.5366
21.1146 | -35.2299 | -31.5439
21.1207 | -35.2304 | -31.5502
21.1268 | -35.231 | -31.5565
21.133 | -35.2315 | -31.5636
21.1391 | -35.232 | -31.5698
21.1452 | -35.2326 | -31.5763
21.1513 | -35.2331 | -31.583
21.1574 | -35.2337 | -31.5898
21.1635 | -35.2341 | -31.5955
21.1696 | -35.2345 | -31.6021
21.1757 | -35.235 | -31.6088
21.1818 | -35.2355 | -31.6146
21.1879 | -35.236 | -31.6208
21.194 | -35.2366 | -31.6277
21.2001 | -35.2371 | -31.6343
21.2062 | -35.2375 | -31.6408
21.2123 | -35.2381 | -31.6473
21.2184 | -35.2386 | -31.6548
21.2245 | -35.2391 | -31.6616
21.2306 | -35.2398 | -31.6683
21.2367 | -35.2403 | -31.6755
21.2428 | -35.2408 | -31.6817
21.2489 | -35.2413 | -31.6879
21.255 | -35.2419 | -31.6945
21.2611 | -35.2424 | -31.7016
21.2672 | -35.2429 | -31.7079
21.2733 | -35.2434 | -31.7146
21.2795 | -35.2438 | -31.7206
21.2856 | -35.2444 | -31.7266
21.2917 | -35.2449 | -31.7341
21.2978 | -35.2455 | -31.7407
21.3039 | -35.246 | -31.7485
21.31 | -35.2465 | -31.7552
21.3161 | -35.2469 | -31.7614
21.3222 | -35.2475 | -31.768
21.3283 | -35.248 | -31.7748
21.3344 | -35.2484 | -31.7811
21.3405 | -35.2488 | -31.7879
21.3466 | -35.2494 | -31.7946
21.3527 | -35.25 | -31.8012
21.3588 | -35.2504 | -31.8084
21.3649 | -35.2509 | -31.8149
21.371 | -35.2514 | -31.8208
21.3771 | -35.2518 | -31.8275
21.3832 | -35.2523 | -31.833
21.3893 | -35.2528 | -31.8385
21.3954 | -35.2532 | -31.8438
21.4015 | -35.2535 | -31.8501
21.4076 | -35.254 | -31.8563
21.4137 | -35.2545 | -31.863
21.4199 | -35.2551 | -31.8705
21.426 | -35.2554 | -31.8777
21.4321 | -35.256 | -31.8843
21.4382 | -35.2563 | -31.8902
21.4443 | -35.2569 | -31.897
21.4504 | -35.2572 | -31.904
21.4565 | -35.2574 | -31.9112
21.4626 | -35.2581 | -31.9176
21.4687 | -35.2586 | -31.9246
21.4748 | -35.2591 | -31.931
21.4809 | -35.2595 | -31.9376
21.487 | -35.26 | -31.9438
21.4931 | -35.2607 | -31.9506
21.4992 | -35.2613 | -31.9569
21.5053 | -35.2618 | -31.9636
21.5114 | -35.2622 | -31.9685
21.5175 | -35.2627 | -31.9747
21.5236 | -35.2631 | -31.9813
21.5297 | -35.2633 | -31.988
21.5358 | -35.2639 | -31.995
21.5419 | -35.2644 | -32.0017
21.548 | -35.265 | -32.008
21.5541 | -35.2654 | -32.0142
21.5602 | -35.2658 | -32.0202
21.5664 | -35.2665 | -32.0265
21.5725 | -35.2669 | -32.0331
21.5786 | -35.2674 | -32.0398
21.5847 | -35.2679 | -32.0467
21.5908 | -35.2683 | -32.0538
21.5969 | -35.2688 | -32.0595
21.603 | -35.2692 | -32.0658
21.6091 | -35.2697 | -32.0725
21.6152 | -35.2703 | -32.0786
21.6213 | -35.2708 | -32.0858
21.6274 | -35.2715 | -32.0931
21.6335 | -35.2718 | -32.1006
21.6396 | -35.2724 | -32.1067
21.6457 | -35.273 | -32.113
21.6518 | -35.2735 | -32.1186
21.6579 | -35.274 | -32.1251
21.664 | -35.2744 | -32.1324
21.6701 | -35.2749 | -32.1389
21.6762 | -35.2754 | -32.1445
21.6823 | -35.2758 | -32.1497
21.6884 | -35.2761 | -32.1572
21.6945 | -35.2765 | -32.1638
21.7006 | -35.277 | -32.1703
21.7068 | -35.2774 | -32.1774
21.7129 | -35.278 | -32.1837
21.719 | -35.2785 | -32.1897
21.7251 | -35.279 | -32.196
21.7312 | -35.2796 | -32.2024
21.7373 | -35.2802 | -32.208
21.7434 | -35.2805 | -32.2151
21.7495 | -35.2811 | -32.222
21.7556 | -35.2815 | -32.2286
21.7617 | -35.282 | -32.2361
21.7678 | -35.2827 | -32.2425
21.7739 | -35.283 | -32.2485
21.78 | -35.2835 | -32.2545
21.7861 | -35.284 | -32.2604
21.7922 | -35.2845 | -32.267
21.7983 | -35.2849 | -32.2723
21.8044 | -35.2853 | -32.2786
21.8105 | -35.2858 | -32.2842
21.8166 | -35.2864 | -32.2908
21.8227 | -35.2869 | -32.2972
21.8288 | -35.2874 | -32.3041
21.8349 | -35.2879 | -32.3108
21.841 | -35.2884 | -32.318
21.8471 | -35.2888 | -32.3241
21.8533 | -35.2893 | -32.3301
21.8594 | -35.2899 | -32.3382
21.8655 | -35.2904 | -32.3448
21.8716 | -35.291 | -32.3516
21.8777 | -35.2916 | -32.3576
21.8838 | -35.2921 | -32.3642
21.8899 | -35.2926 | -32.3707
21.896 | -35.2931 | -32.3765
21.9021 | -35.2935 | -32.3834
21.9082 | -35.2939 | -32.3897
21.9143 | -35.2943 | -32.3961
21.9204 | -35.2946 | -32.4035
21.9265 | -35.295 | -32.4104
21.9326 | -35.2955 | -32.4167
21.9387 | -35.2959 | -32.4226
21.9448 | -35.2964 | -32.4292
21.9509 | -35.2967 | -32.4353
21.957 | -35.2973 | -32.4418
21.9631 | -35.2977 | -32.4489
21.9692 | -35.2982 | -32.4549
21.9753 | -35.2987 | -32.4619
21.9814 | -35.2991 | -32.4673
21.9875 | -35.2996 | -32.4735
21.9937 | -35.3 | -32.4802
21.9998 | -35.3005 | -32.4862
22.0059 | -35.3009 | -32.4927
22.012 | -35.3014 | -32.499
22.0181 | -35.3018 | -32.505
22.0242 | -35.3023 | -32.5113
22.0303 | -35.3028 | -32.5178
22.0364 | -35.3034 | -32.5237
22.0425 | -35.3039 | -32.5298
22.0486 | -35.3044 | -32.5364
22.0547 | -35.3047 | -32.5424
22.0608 | -35.3051 | -32.5488
22.0669 | -35.3056 | -32.5556
22.073 | -35.3062 | -32.5617
22.0791 | -35.3066 | -32.5684
22.0852 | -35.3071 | -32.5743
22.0913 | -35.3076 | -32.5815
22.0974 | -35.3081 | -32.5881
22.1035 | -35.3085 | -32.5944
22.1096 | -35.3089 | -32.6016
22.1157 | -35.3094 | -32.608
22.1218 | -35.3099 | -32.6141
22.1279 | -35.3104 | -32.6196
22.134 | -35.3109 | -32.6257
22.1402 | -35.3113 | -32.6317
22.1463 | -35.3117 | -32.638
22.1524 | -35.3123 | -32.6448
22.1585 | -35.3127 | -32.6514
22.1646 | -35.3131 | -32.6569
22.1707 | -35.3136 | -32.6636
22.1768 | -35.3141 | -32.6701
22.1829 | -35.3145 | -32.6764
22.189 | -35.315 | -32.683
22.1951 | -35.3155 | -32.6891
22.2012 | -35.316 | -32.696
22.2073 | -35.3164 | -32.7025
22.2134 | -35.317 | -32.7088
22.2195 | -35.3175 | -32.7155
22.2256 | -35.3179 | -32.7219
22.2317 | -35.3183 | -32.7274
22.2378 | -35.3188 | -32.7333
22.2439 | -35.3193 | -32.7399
22.25 | -35.3198 | -32.7459
22.2561 | -35.3203 | -32.7525
22.2622 | -35.3208 | -32.7591
22.2683 | -35.3213 | -32.7654
22.2744 | -35.3218 | -32.772
22.2806 | -35.3223 | -32.7779
22.2867 | -35.3227 | -32.7844
22.2928 | -35.3232 | -32.7908
22.2989 | -35.3237 | -32.7972
22.305 | -35.3242 | -32.8027
22.3111 | -35.3247 | -32.8092
22.3172 | -35.3252 | -32.8156
22.3233 | -35.3257 | -32.8216
22.3294 | -35.3262 | -32.8275
22.3355 | -35.3267 | -32.833
22.3416 | -35.3272 | -32.839
22.3477 | -35.3276 | -32.8451
22.3538 | -35.3281 | -32.8509
22.3599 | -35.3286 | -32.8574
22.366 | -35.329 | -32.8637
22.3721 | -35.3295 | -32.8693
22.3782 | -35.33 | -32.8756
22.3843 | -35.3304 | -32.882
22.3904 | -35.3309 | -32.8888
22.3965 | -35.3313 | -32.8953
22.4026 | -35.3318 | -32.9016
22.4087 | -35.3322 | -32.9074
22.4148 | -35.3326 | -32.9141
22.4209 | -35.333 | -32.9202
22.4271 | -35.3335 | -32.9261
22.4332 | -35.334 | -32.9321
22.4393 | -35.3344 | -32.9382
22.4454 | -35.3348 | -32.9441
22.4515 | -35.3352 | -32.9511
22.4576 | -35.3357 | -32.9571
22.4637 | -35.3361 | -32.9636
22.4698 | -35.3366 | -32.9703
22.4759 | -35.337 | -32.9762
22.482 | -35.3375 | -32.9823
22.4881 | -35.338 | -32.9888
22.4942 | -35.3384 | -32.9953
22.5003 | -35.3389 | -33.0018
22.5064 | -35.3392 | -33.0079
22.5125 | -35.3397 | -33.0148
22.5186 | -35.3402 | -33.0206
22.5247 | -35.3406 | -33.0269
22.5308 | -35.3411 | -33.0331
22.5369 | -35.3415 | -33.0401
22.543 | -35.342 | -33.0456
22.5491 | -35.3424 | -33.0515
22.5552 | -35.3428 | -33.0579
22.5613 | -35.3433 | -33.0643
22.5675 | -35.3437 | -33.0715
22.5736 | -35.3441 | -33.078
22.5797 | -35.3446 | -33.0848
22.5858 | -35.3451 | -33.0911
22.5919 | -35.3455 | -33.0975
22.598 | -35.346 | -33.1038
22.6041 | -35.3464 | -33.1098
22.6102 | -35.3467 | -33.1156
22.6163 | -35.3472 | -33.1217
22.6224 | -35.3476 | -33.1278
22.6285 | -35.348 | -33.1343
22.6346 | -35.3485 | -33.1405
22.6407 | -35.3489 | -33.1468
22.6468 | -35.3494 | -33.153
22.6529 | -35.3497 | -33.1593
22.659 | -35.3502 | -33.1657
22.6651 | -35.3505 | -33.1722
22.6712 | -35.351 | -33.1783
22.6773 | -35.3514 | -33.184
22.6834 | -35.3519 | -33.1906
22.6895 | -35.3523 | -33.197
22.6956 | -35.3528 | -33.2032
22.7017 | -35.3533 | -33.2095
22.7079 | -35.3537 | -33.2155
22.714 | -35.3542 | -33.2216
22.7201 | -35.3547 | -33.2278
22.7262 | -35.3552 | -33.2338
22.7323 | -35.3556 | -33.2401
22.7384 | -35.3561 | -33.246
22.7445 | -35.3565 | -33.2526
22.7506 | -35.357 | -33.2589
22.7567 | -35.3574 | -33.2654
22.7628 | -35.3579 | -33.2716
22.7689 | -35.3584 | -33.2775
22.775 | -35.3588 | -33.2837
22.7811 | -35.3593 | -33.2899
22.7872 | -35.3597 | -33.2963
22.7933 | -35.3603 | -33.3022
22.7994 | -35.3607 | -33.3082
22.8055 | -35.3612 | -33.3146
22.8116 | -35.3617 | -33.3213
22.8177 | -35.362 | -33.3277
22.8238 | -35.3625 | -33.3338
22.8299 | -35.3629 | -33.3397
22.836 | -35.3633 | -33.346
22.8421 | -35.3638 | -33.3522
22.8482 | -35.3643 | -33.3584
22.8544 | -35.3647 | -33.3647
22.8605 | -35.3652 | -33.3705
22.8666 | -35.3656 | -33.3766
22.8727 | -35.366 | -33.383
22.8788 | -35.3664 | -33.3893
22.8849 | -35.3669 | -33.3954
22.891 | -35.3674 | -33.401
22.8971 | -35.3678 | -33.4068
22.9032 | -35.3683 | -33.4129
22.9093 | -35.3687 | -33.4197
22.9154 | -35.3691 | -33.4261
22.9215 | -35.3696 | -33.4323
22.9276 | -35.37 | -33.4381
22.9337 | -35.3704 | -33.4439
22.9398 | -35.3708 | -33.4506
22.9459 | -35.3713 | -33.4566
22.952 | -35.3717 | -33.4625
22.9581 | -35.3722 | -33.469
22.9642 | -35.3726 | -33.4753
22.9703 | -35.373 | -33.4808
22.9764 | -35.3735 | -33.4872
22.9825 | -35.374 | -33.493
22.9886 | -35.3744 | -33.499
22.9948 | -35.3748 | -33.5059
23.0009 | -35.3752 | -33.5123
23.007 | -35.3756 | -33.519
23.0131 | -35.3761 | -33.525
23.0192 | -35.3765 | -33.5309
23.0253 | -35.3769 | -33.5368
23.0314 | -35.3774 | -33.543
23.0375 | -35.3779 | -33.5495
23.0436 | -35.3783 | -33.5552
23.0497 | -35.3787 | -33.5619
23.0558 | -35.3791 | -33.5679
23.0619 | -35.3796 | -33.574
23.068 | -35.3801 | -33.5801
23.0741 | -35.3805 | -33.5861
23.0802 | -35.381 | -33.5929
23.0863 | -35.3814 | -33.5994
23.0924 | -35.3819 | -33.6053
23.0985 | -35.3823 | -33.6111
23.1046 | -35.3827 | -33.6172
23.1107 | -35.3831 | -33.6235
23.1168 | -35.3834 | -33.6296
23.1229 | -35.3839 | -33.6359
23.129 | -35.3843 | -33.6419
23.1351 | -35.3847 | -33.6481
23.1413 | -35.3852 | -33.6541
23.1474 | -35.3857 | -33.6604
23.1535 | -35.3861 | -33.6666
23.1596 | -35.3866 | -33.6727
23.1657 | -35.387 | -33.6785
23.1718 | -35.3875 | -33.6853
23.1779 | -35.3878 | -33.6914
23.184 | -35.3883 | -33.6971
23.1901 | -35.3888 | -33.7033
23.1962 | -35.3891 | -33.7092
23.2023 | -35.3895 | -33.715
23.2084 | -35.39 | -33.7212
23.2145 | -35.3905 | -33.7276
23.2206 | -35.3909 | -33.7338
23.2267 | -35.3913 | -33.74
23.2328 | -35.3917 | -33.7463
23.2389 | -35.3921 | -33.753
23.245 | -35.3926 | -33.7591
23.2511 | -35.393 | -33.7648
23.2572 | -35.3934 | -33.7712
23.2633 | -35.3937 | -33.777
23.2694 | -35.3942 | -33.7832
23.2755 | -35.3946 | -33.7893
23.2817 | -35.395 | -33.7952
23.2878 | -35.3954 | -33.8013
23.2939 | -35.3959 | -33.8078
23.3 | -35.3963 | -33.814
23.3061 | -35.3968 | -33.8198
23.3122 | -35.3972 | -33.826
23.3183 | -35.3976 | -33.8319
23.3244 | -35.3979 | -33.8378
23.3305 | -35.3983 | -33.8441
23.3366 | -35.3987 | -33.8509
23.3427 | -35.3992 | -33.8572
23.3488 | -35.3996 | -33.8633
23.3549 | -35.4 | -33.8699
23.361 | -35.4006 | -33.8761
23.3671 | -35.401 | -33.8824
23.3732 | -35.4015 | -33.8884
23.3793 | -35.4019 | -33.8943
23.3854 | -35.4023 | -33.9002
23.3915 | -35.4027 | -33.9058
23.3976 | -35.4032 | -33.9121
23.4037 | -35.4036 | -33.9184
23.4098 | -35.404 | -33.9245
23.4159 | -35.4044 | -33.9303
23.422 | -35.4049 | -33.9363
23.4282 | -35.4053 | -33.9416
23.4343 | -35.4057 | -33.9476
23.4404 | -35.4061 | -33.9539
23.4465 | -35.4065 | -33.9598
23.4526 | -35.407 | -33.9663
23.4587 | -35.4074 | -33.9723
23.4648 | -35.4078 | -33.9782
23.4709 | -35.4082 | -33.9842
23.477 | -35.4086 | -33.9904
23.4831 | -35.409 | -33.9964
23.4892 | -35.4096 | -34.0026
23.4953 | -35.4099 | -34.0083
23.5014 | -35.4103 | -34.0137
23.5075 | -35.4107 | -34.0193
23.5136 | -35.4111 | -34.0253
23.5197 | -35.4115 | -34.0315
23.5258 | -35.4119 | -34.0375
23.5319 | -35.4123 | -34.043
23.538 | -35.4127 | -34.0486
23.5441 | -35.4132 | -34.0552
23.5502 | -35.4136 | -34.0612
23.5563 | -35.4141 | -34.0667
23.5624 | -35.4144 | -34.0732
23.5686 | -35.4149 | -34.0789
23.5747 | -35.4153 | -34.0845
23.5808 | -35.4157 | -34.0903
23.5869 | -35.4162 | -34.0961
23.593 | -35.4166 | -34.1019
23.5991 | -35.417 | -34.1084
23.6052 | -35.4175 | -34.114
23.6113 | -35.4178 | -34.12
23.6174 | -35.4182 | -34.1257
23.6235 | -35.4187 | -34.1318
23.6296 | -35.4191 | -34.1382
23.6357 | -35.4195 | -34.1449
23.6418 | -35.42 | -34.151
23.6479 | -35.4204 | -34.157
23.654 | -35.4208 | -34.1629
23.6601 | -35.4212 | -34.1687
23.6662 | -35.4217 | -34.1749
23.6723 | -35.422 | -34.1812
23.6784 | -35.4225 | -34.1872
23.6845 | -35.4229 | -34.1932
23.6906 | -35.4233 | -34.1991
23.6967 | -35.4237 | -34.2044
23.7028 | -35.424 | -34.2105
23.7089 | -35.4243 | -34.2156
23.7151 | -35.4248 | -34.2216
23.7212 | -35.4252 | -34.2274
23.7273 | -35.4256 | -34.2336
23.7334 | -35.426 | -34.2393
23.7395 | -35.4264 | -34.2444
23.7456 | -35.4267 | -34.2509
23.7517 | -35.427 | -34.2567
23.7578 | -35.4274 | -34.2627
23.7639 | -35.4278 | -34.2682
23.77 | -35.4283 | -34.2745
23.7761 | -35.4286 | -34.2804
23.7822 | -35.4291 | -34.2855
23.7883 | -35.4295 | -34.2913
23.7944 | -35.43 | -34.2973
23.8005 | -35.4304 | -34.3034
23.8066 | -35.4308 | -34.3093
23.8127 | -35.4313 | -34.3151
23.8188 | -35.4316 | -34.3214
23.8249 | -35.4321 | -34.3274
23.831 | -35.4325 | -34.3335
23.8371 | -35.4329 | -34.3391
23.8432 | -35.4334 | -34.3448
23.8493 | -35.4338 | -34.3507
23.8555 | -35.4343 | -34.3569
23.8616 | -35.4346 | -34.363
23.8677 | -35.435 | -34.3689
23.8738 | -35.4354 | -34.3752
23.8799 | -35.4359 | -34.3806
23.886 | -35.4363 | -34.3859
23.8921 | -35.4366 | -34.3919
23.8982 | -35.4371 | -34.3976
23.9043 | -35.4374 | -34.4036
23.9104 | -35.4378 | -34.4094
23.9165 | -35.4383 | -34.415
23.9226 | -35.4388 | -34.4208
23.9287 | -35.4393 | -34.4267
23.9348 | -35.4396 | -34.4326
23.9409 | -35.4401 | -34.4389
23.947 | -35.4405 | -34.4441
23.9531 | -35.441 | -34.4503
23.9592 | -35.4414 | -34.4562
23.9653 | -35.4418 | -34.4622
23.9714 | -35.4422 | -34.4684
23.9775 | -35.4426 | -34.4739
23.9836 | -35.443 | -34.4795
23.9897 | -35.4435 | -34.4852
23.9958 | -35.4439 | -34.491
24.002 | -35.4444 | -34.497
24.0081 | -35.4448 | -34.503
24.0142 | -35.4452 | -34.5095
24.0203 | -35.4456 | -34.515
24.0264 | -35.4461 | -34.5212
24.0325 | -35.4465 | -34.5273
24.0386 | -35.4469 | -34.5331
24.0447 | -35.4473 | -34.539
24.0508 | -35.4477 | -34.5446
24.0569 | -35.448 | -34.5505
24.063 | -35.4485 | -34.5561
24.0691 | -35.4489 | -34.562
24.0752 | -35.4494 | -34.568
24.0813 | -35.4498 | -34.5741
24.0874 | -35.4503 | -34.5798
24.0935 | -35.4507 | -34.5853
24.0996 | -35.451 | -34.5908
24.1057 | -35.4514 | -34.5959
24.1118 | -35.4518 | -34.602
24.1179 | -35.4522 | -34.6084
24.124 | -35.4526 | -34.614
24.1301 | -35.453 | -34.6202
24.1362 | -35.4534 | -34.6262
24.1424 | -35.4538 | -34.6322
24.1485 | -35.4542 | -34.6382
24.1546 | -35.4545 | -34.6443
24.1607 | -35.4549 | -34.6503
24.1668 | -35.4553 | -34.6563
24.1729 | -35.4557 | -34.6626
24.179 | -35.4562 | -34.6689
24.1851 | -35.4566 | -34.6745
24.1912 | -35.457 | -34.68
24.1973 | -35.4575 | -34.6861
24.2034 | -35.4579 | -34.6918
24.2095 | -35.4583 | -34.6976
24.2156 | -35.4587 | -34.7039
24.2217 | -35.4592 | -34.7099
24.2278 | -35.4596 | -34.7162
24.2339 | -35.4599 | -34.722
24.24 | -35.4603 | -34.7278
24.2461 | -35.4607 | -34.7336
24.2522 | -35.4612 | -34.7395
24.2583 | -35.4616 | -34.7457
24.2644 | -35.462 | -34.7518
24.2705 | -35.4625 | -34.7576
24.2766 | -35.463 | -34.7631
24.2827 | -35.4634 | -34.7689
24.2889 | -35.4638 | -34.7748
24.295 | -35.4642 | -34.7803
24.3011 | -35.4647 | -34.7864
24.3072 | -35.4651 | -34.7928
24.3133 | -35.4654 | -34.7991
24.3194 | -35.4659 | -34.8051
24.3255 | -35.4663 | -34.8112
24.3316 | -35.4667 | -34.8173
24.3377 | -35.467 | -34.8232
24.3438 | -35.4674 | -34.8296
24.3499 | -35.4677 | -34.8354
24.356 | -35.468 | -34.8411
24.3621 | -35.4683 | -34.8469
24.3682 | -35.4688 | -34.8528
24.3743 | -35.4693 | -34.8592
24.3804 | -35.4697 | -34.8645
24.3865 | -35.4701 | -34.8705
24.3926 | -35.4705 | -34.8764
24.3987 | -35.4708 | -34.8819
24.4048 | -35.4712 | -34.8881
24.4109 | -35.4714 | -34.8935
24.417 | -35.4718 | -34.8991
24.4231 | -35.4721 | -34.9052
24.4293 | -35.4725 | -34.9107
24.4354 | -35.473 | -34.9159
24.4415 | -35.4735 | -34.9217
24.4476 | -35.4738 | -34.9266
24.4537 | -35.4743 | -34.9323
24.4598 | -35.4748 | -34.9384
24.4659 | -35.4752 | -34.945
24.472 | -35.4756 | -34.9505
24.4781 | -35.476 | -34.9564
24.4842 | -35.4763 | -34.9618
24.4903 | -35.4768 | -34.9677
24.4964 | -35.4772 | -34.9733
24.5025 | -35.4775 | -34.9797
24.5086 | -35.478 | -34.9853
24.5147 | -35.4783 | -34.9912
24.5208 | -35.4787 | -34.9963
24.5269 | -35.4792 | -35.0017
24.533 | -35.4796 | -35.0072
24.5391 | -35.4801 | -35.013
24.5452 | -35.4804 | -35.0185
24.5513 | -35.4808 | -35.0238
24.5574 | -35.4812 | -35.0298
24.5635 | -35.4816 | -35.0353
24.5696 | -35.482 | -35.0413
24.5758 | -35.4824 | -35.0472
24.5819 | -35.4828 | -35.0526
24.588 | -35.4832 | -35.0583
24.5941 | -35.4836 | -35.0642
24.6002 | -35.484 | -35.0699
24.6063 | -35.4844 | -35.0753
24.6124 | -35.4848 | -35.0806
24.6185 | -35.4852 | -35.0863
24.6246 | -35.4856 | -35.0919
24.6307 | -35.486 | -35.0978
24.6368 | -35.4864 | -35.1035
24.6429 | -35.4867 | -35.1094
24.649 | -35.4871 | -35.1153
24.6551 | -35.4875 | -35.1212
24.6612 | -35.488 | -35.1275
24.6673 | -35.4884 | -35.1331
24.6734 | -35.4887 | -35.139
24.6795 | -35.4891 | -35.1446
24.6856 | -35.4896 | -35.1507
24.6917 | -35.49 | -35.1564
24.6978 | -35.4903 | -35.1617
24.7039 | -35.4908 | -35.1677
24.71 | -35.4912 | -35.1733
24.7162 | -35.4916 | -35.1798
24.7223 | -35.492 | -35.1853
24.7284 | -35.4924 | -35.1906
24.7345 | -35.4929 | -35.1968
24.7406 | -35.4933 | -35.2028
24.7467 | -35.4937 | -35.2086
24.7528 | -35.4941 | -35.2146
24.7589 | -35.4944 | -35.2204
24.765 | -35.4947 | -35.2259
24.7711 | -35.495 | -35.2319
24.7772 | -35.4954 | -35.2376
24.7833 | -35.4958 | -35.2427
24.7894 | -35.4963 | -35.2484
24.7955 | -35.4967 | -35.2542
24.8016 | -35.497 | -35.2599
24.8077 | -35.4974 | -35.2666
24.8138 | -35.4977 | -35.2723
24.8199 | -35.4982 | -35.278
24.826 | -35.4985 | -35.2834
24.8321 | -35.4989 | -35.2886
24.8382 | -35.4993 | -35.2944
24.8443 | -35.4997 | -35.3002
24.8504 | -35.5001 | -35.306
24.8565 | -35.5006 | -35.3119
24.8627 | -35.501 | -35.3175
24.8688 | -35.5013 | -35.3232
24.8749 | -35.5017 | -35.3287
24.881 | -35.502 | -35.3345
24.8871 | -35.5024 | -35.3401
24.8932 | -35.5029 | -35.346
24.8993 | -35.5032 | -35.3514
24.9054 | -35.5037 | -35.3561
24.9115 | -35.5041 | -35.3621
24.9176 | -35.5046 | -35.3685
24.9237 | -35.505 | -35.3742
24.9298 | -35.5054 | -35.38
24.9359 | -35.5057 | -35.386
24.942 | -35.5061 | -35.3917
24.9481 | -35.5064 | -35.3976
24.9542 | -35.5068 | -35.4039
24.9603 | -35.5072 | -35.4094
24.9664 | -35.5076 | -35.4144
24.9725 | -35.5079 | -35.4206
24.9786 | -35.5083 | -35.425
24.9847 | -35.5087 | -35.4304
24.9908 | -35.509 | -35.4363
24.9969 | -35.5093 | -35.4416
25.0031 | -35.5096 | -35.448
25.0092 | -35.51 | -35.4531
25.0153 | -35.5105 | -35.4584
25.0214 | -35.5108 | -35.4646
25.0275 | -35.5112 | -35.4702
25.0336 | -35.5116 | -35.4767
25.0397 | -35.5118 | -35.4825
25.0458 | -35.5121 | -35.4883
25.0519 | -35.5125 | -35.4944
25.058 | -35.5129 | -35.5001
25.0641 | -35.5133 | -35.5059
25.0702 | -35.5137 | -35.5115
25.0763 | -35.514 | -35.5175
25.0824 | -35.5143 | -35.5234
25.0885 | -35.5147 | -35.5289
25.0946 | -35.5151 | -35.5347
25.1007 | -35.5154 | -35.5406
25.1068 | -35.5159 | -35.5458
25.1129 | -35.5163 | -35.5511
25.119 | -35.5167 | -35.5563
25.1251 | -35.5172 | -35.5614
25.1312 | -35.5176 | -35.5671
25.1373 | -35.5181 | -35.5732
25.1435 | -35.5186 | -35.579
25.1496 | -35.5189 | -35.5838
25.1557 | -35.5193 | -35.5891
25.1618 | -35.5196 | -35.5948
25.1679 | -35.52 | -35.5999
25.174 | -35.5204 | -35.6054
25.1801 | -35.5208 | -35.6115
25.1862 | -35.5213 | -35.6168
25.1923 | -35.5216 | -35.6226
25.1984 | -35.5221 | -35.6276
25.2045 | -35.5225 | -35.6338
25.2106 | -35.5229 | -35.6397
25.2167 | -35.5233 | -35.6451
25.2228 | -35.5236 | -35.6508
25.2289 | -35.524 | -35.6564
25.235 | -35.5245 | -35.6623
25.2411 | -35.5249 | -35.6678
25.2472 | -35.5253 | -35.6736
25.2533 | -35.5257 | -35.6791
25.2594 | -35.5259 | -35.6849
25.2655 | -35.5263 | -35.6905
25.2716 | -35.5268 | -35.6959
25.2777 | -35.5273 | -35.7016
25.2838 | -35.5276 | -35.7072
25.29 | -35.528 | -35.7129
25.2961 | -35.5284 | -35.7191
25.3022 | -35.5288 | -35.7246
25.3083 | -35.5292 | -35.73
25.3144 | -35.5295 | -35.7357
25.3205 | -35.5299 | -35.7412
25.3266 | -35.5303 | -35.747
25.3327 | -35.5306 | -35.7519
25.3388 | -35.5309 | -35.7573
25.3449 | -35.5313 | -35.7635
25.351 | -35.5317 | -35.7692
25.3571 | -35.5321 | -35.7753
25.3632 | -35.5324 | -35.7805
25.3693 | -35.5327 | -35.7862
25.3754 | -35.5331 | -35.792
25.3815 | -35.5335 | -35.7974
25.3876 | -35.5339 | -35.8031
25.3937 | -35.5343 | -35.8087
25.3998 | -35.5348 | -35.8144
25.4059 | -35.5352 | -35.8202
25.412 | -35.5356 | -35.8253
25.4181 | -35.536 | -35.831
25.4242 | -35.5364 | -35.837
25.4304 | -35.5367 | -35.8424
25.4365 | -35.537 | -35.848
25.4426 | -35.5372 | -35.8532
25.4487 | -35.5376 | -35.8586
25.4548 | -35.538 | -35.8637
25.4609 | -35.5384 | -35.8688
25.467 | -35.5386 | -35.874
25.4731 | -35.539 | -35.8799
25.4792 | -35.5395 | -35.8849
25.4853 | -35.5399 | -35.8909
25.4914 | -35.5402 | -35.8962
25.4975 | -35.5406 | -35.9021
25.5036 | -35.5408 | -35.9083
25.5097 | -35.5412 | -35.9143
25.5158 | -35.5414 | -35.9198
25.5219 | -35.5418 | -35.9249
25.528 | -35.5422 | -35.9298
25.5341 | -35.5424 | -35.9354
25.5402 | -35.5428 | -35.9413
25.5463 | -35.5431 | -35.9473
25.5524 | -35.5434 | -35.9528
25.5585 | -35.5438 | -35.9578
25.5646 | -35.5442 | -35.9635
25.5707 | -35.5445 | -35.9678
25.5769 | -35.5448 | -35.9742
25.583 | -35.5452 | -35.9802
25.5891 | -35.5457 | -35.9851
25.5952 | -35.5461 | -35.9905
25.6013 | -35.5465 | -35.9967
25.6074 | -35.5469 | -36.0027
25.6135 | -35.5473 | -36.0087
25.6196 | -35.5477 | -36.0148
25.6257 | -35.548 | -36.0205
25.6318 | -35.5484 | -36.0262
25.6379 | -35.5488 | -36.0319
25.644 | -35.5492 | -36.0368
25.6501 | -35.5496 | -36.0419
25.6562 | -35.55 | -36.048
25.6623 | -35.5504 | -36.0538
25.6684 | -35.5507 | -36.0591
25.6745 | -35.551 | -36.0647
25.6806 | -35.5515 | -36.0697
25.6867 | -35.5518 | -36.0755
25.6928 | -35.5521 | -36.0805
25.6989 | -35.5524 | -36.0858
25.705 | -35.5528 | -36.0919
25.7111 | -35.5532 | -36.0979
25.7173 | -35.5535 | -36.1024
25.7234 | -35.5539 | -36.108
25.7295 | -35.5543 | -36.1134
25.7356 | -35.5547 | -36.1191
25.7417 | -35.5551 | -36.1247
25.7478 | -35.5553 | -36.1303
25.7539 | -35.5557 | -36.1355
25.76 | -35.5561 | -36.1408
25.7661 | -35.5565 | -36.1463
25.7722 | -35.5569 | -36.152
25.7783 | -35.5573 | -36.1572
25.7844 | -35.5577 | -36.1625
25.7905 | -35.558 | -36.1686
25.7966 | -35.5583 | -36.1738
25.8027 | -35.5586 | -36.1793
25.8088 | -35.559 | -36.1853
25.8149 | -35.5594 | -36.1909
25.821 | -35.5597 | -36.1961
25.8271 | -35.5601 | -36.2017
25.8332 | -35.5605 | -36.208
25.8393 | -35.5607 | -36.2131
25.8454 | -35.5612 | -36.2181
25.8515 | -35.5615 | -36.2233
25.8576 | -35.5619 | -36.2283
25.8638 | -35.5623 | -36.234
25.8699 | -35.5627 | -36.2393
25.876 | -35.563 | -36.2451
25.8821 | -35.5634 | -36.2511
25.8882 | -35.5636 | -36.2564
25.8943 | -35.564 | -36.2621
25.9004 | -35.5643 | -36.2678
25.9065 | -35.5647 | -36.2733
25.9126 | -35.5651 | -36.2787
25.9187 | -35.5653 | -36.2845
25.9248 | -35.5658 | -36.2899
25.9309 | -35.5661 | -36.2958
25.937 | -35.5665 | -36.3009
25.9431 | -35.5669 | -36.3064
25.9492 | -35.5672 | -36.3113
25.9553 | -35.5675 | -36.3166
25.9614 | -35.5678 | -36.3217
25.9675 | -35.5682 | -36.3271
25.9736 | -35.5685 | -36.3321
25.9797 | -35.5687 | -36.3375
25.9858 | -35.569 | -36.3426
25.9919 | -35.5693 | -36.3485
25.998 | -35.5697 | -36.3536
26.0042 | -35.5701 | -36.3594
26.0103 | -35.5706 | -36.3646
26.0164 | -35.5709 | -36.3698
26.0225 | -35.5712 | -36.3755
26.0286 | -35.5715 | -36.3811
26.0347 | -35.5717 | -36.3873
26.0408 | -35.5721 | -36.3928
26.0469 | -35.5725 | -36.3981
26.053 | -35.5729 | -36.4037
26.0591 | -35.5734 | -36.4098
26.0652 | -35.5738 | -36.4159
26.0713 | -35.5741 | -36.4218
26.0774 | -35.5744 | -36.4274
26.0835 | -35.5747 | -36.4329
26.0896 | -35.5751 | -36.4381
26.0957 | -35.5756 | -36.4433
26.1018 | -35.5758 | -36.4488
26.1079 | -35.5762 | -36.4544
26.114 | -35.5764 | -36.4605
26.1201 | -35.5767 | -36.4665
26.1262 | -35.577 | -36.4725
26.1323 | -35.5773 | -36.4775
26.1384 | -35.5776 | -36.4824
26.1445 | -35.5779 | -36.4887
26.1507 | -35.5782 | -36.4944
26.1568 | -35.5786 | -36.5008
26.1629 | -35.5789 | -36.5059
26.169 | -35.5792 | -36.5108
26.1751 | -35.5795 | -36.5158
26.1812 | -35.5799 | -36.5209
26.1873 | -35.5803 | -36.5267
26.1934 | -35.5808 | -36.5327
26.1995 | -35.581 | -36.5382
26.2056 | -35.5813 | -36.5434
26.2117 | -35.5817 | -36.5486
26.2178 | -35.582 | -36.5543
26.2239 | -35.5825 | -36.5597
26.23 | -35.5827 | -36.5648
26.2361 | -35.583 | -36.5701
26.2422 | -35.5833 | -36.5753
26.2483 | -35.5836 | -36.5811
26.2544 | -35.584 | -36.5862
26.2605 | -35.5843 | -36.592
26.2666 | -35.5846 | -36.5974
26.2727 | -35.5849 | -36.6031
26.2788 | -35.5854 | -36.6077
26.2849 | -35.5857 | -36.6132
26.2911 | -35.586 | -36.6188
26.2972 | -35.5865 | -36.625
26.3033 | -35.587 | -36.6305
26.3094 | -35.5872 | -36.6361
26.3155 | -35.5875 | -36.6422
26.3216 | -35.5879 | -36.6473
26.3277 | -35.5883 | -36.6528
26.3338 | -35.5886 | -36.6577
26.3399 | -35.5889 | -36.6622
26.346 | -35.5893 | -36.667
26.3521 | -35.5895 | -36.673
26.3582 | -35.59 | -36.6789
26.3643 | -35.5903 | -36.6846
26.3704 | -35.5908 | -36.6903
26.3765 | -35.5912 | -36.696
26.3826 | -35.5915 | -36.7014
26.3887 | -35.5919 | -36.7072
26.3948 | -35.5923 | -36.7122
26.4009 | -35.5927 | -36.718
26.407 | -35.593 | -36.7228
26.4131 | -35.5934 | -36.7283
26.4192 | -35.5938 | -36.734
26.4253 | -35.5941 | -36.7384
26.4314 | -35.5945 | -36.7442
26.4376 | -35.5949 | -36.7499
26.4437 | -35.5953 | -36.7552
26.4498 | -35.5958 | -36.7603
26.4559 | -35.5961 | -36.7654
26.462 | -35.5965 | -36.7707
26.4681 | -35.5968 | -36.7761
26.4742 | -35.5971 | -36.7815
26.4803 | -35.5974 | -36.7871
26.4864 | -35.5977 | -36.7927
26.4925 | -35.598 | -36.7983
26.4986 | -35.5983 | -36.8029
26.5047 | -35.5986 | -36.8084
26.5108 | -35.5989 | -36.8139
26.5169 | -35.5992 | -36.8193
26.523 | -35.5995 | -36.825
26.5291 | -35.5999 | -36.8305
26.5352 | -35.6002 | -36.8361
26.5413 | -35.6006 | -36.8415
26.5474 | -35.601 | -36.8477
26.5535 | -35.6013 | -36.8535
26.5596 | -35.6016 | -36.8604
26.5657 | -35.6021 | -36.8659
26.5718 | -35.6025 | -36.8711
26.578 | -35.6028 | -36.8769
26.5841 | -35.6033 | -36.8823
26.5902 | -35.6035 | -36.8879
26.5963 | -35.6038 | -36.8944
26.6024 | -35.6042 | -36.8993
26.6085 | -35.6045 | -36.904
26.6146 | -35.6049 | -36.9095
26.6207 | -35.6052 | -36.9157
26.6268 | -35.6055 | -36.9216
26.6329 | -35.6058 | -36.927
26.639 | -35.6061 | -36.9326
26.6451 | -35.6065 | -36.9375
26.6512 | -35.6068 | -36.9435
26.6573 | -35.6071 | -36.9488
26.6634 | -35.6072 | -36.9537
26.6695 | -35.6075 | -36.9597
26.6756 | -35.608 | -36.9658
26.6817 | -35.6083 | -36.9716
26.6878 | -35.6088 | -36.977
26.6939 | -35.609 | -36.9816
26.7 | -35.6094 | -36.987
26.7061 | -35.6098 | -36.9924
26.7122 | -35.6102 | -36.9979
26.7183 | -35.6105 | -37.0028
26.7245 | -35.6107 | -37.0086
26.7306 | -35.611 | -37.0141
26.7367 | -35.6113 | -37.0192
26.7428 | -35.6118 | -37.0247
26.7489 | -35.6121 | -37.031
26.755 | -35.6125 | -37.0364
26.7611 | -35.6129 | -37.0421
26.7672 | -35.6133 | -37.0472
26.7733 | -35.6136 | -37.0514
26.7794 | -35.6139 | -37.057
26.7855 | -35.6143 | -37.0621
26.7916 | -35.6145 | -37.0679
26.7977 | -35.6148 | -37.0734
26.8038 | -35.6151 | -37.079
26.8099 | -35.6153 | -37.0848
26.816 | -35.6156 | -37.091
26.8221 | -35.6159 | -37.0968
26.8282 | -35.6163 | -37.1019
26.8343 | -35.6166 | -37.1076
26.8404 | -35.6169 | -37.1139
26.8465 | -35.6171 | -37.1194
26.8526 | -35.6175 | -37.1251
26.8587 | -35.6178 | -37.1311
26.8649 | -35.6181 | -37.1357
26.871 | -35.6184 | -37.1415
26.8771 | -35.6187 | -37.1472
26.8832 | -35.619 | -37.1527
26.8893 | -35.6193 | -37.1589
26.8954 | -35.6195 | -37.164
26.9015 | -35.6198 | -37.1695
26.9076 | -35.6202 | -37.1757
26.9137 | -35.6206 | -37.1816
26.9198 | -35.6211 | -37.187
26.9259 | -35.6214 | -37.192
26.932 | -35.6219 | -37.1964
26.9381 | -35.6223 | -37.2025
26.9442 | -35.6226 | -37.2074
26.9503 | -35.6229 | -37.2124
26.9564 | -35.6233 | -37.218
26.9625 | -35.6235 | -37.2222
26.9686 | -35.6239 | -37.2277
26.9747 | -35.6242 | -37.2333
26.9808 | -35.6246 | -37.2392
26.9869 | -35.6249 | -37.2443
26.993 | -35.6252 | -37.2504
26.9991 | -35.6255 | -37.2565
27.0052 | -35.6257 | -37.2621
27.0114 | -35.6259 | -37.2682
27.0175 | -35.6261 | -37.2738
27.0236 | -35.6264 | -37.2801
27.0297 | -35.6267 | -37.2858
27.0358 | -35.6269 | -37.2907
27.0419 | -35.6273 | -37.2964
27.048 | -35.6277 | -37.3014
27.0541 | -35.6278 | -37.3076
27.0602 | -35.6281 | -37.3121
27.0663 | -35.6284 | -37.3169
27.0724 | -35.6288 | -37.322
27.0785 | -35.6293 | -37.3273
27.0846 | -35.6297 | -37.3328
27.0907 | -35.6301 | -37.3387
27.0968 | -35.6302 | -37.3447
27.1029 | -35.6305 | -37.3498
27.109 | -35.6308 | -37.356
27.1151 | -35.6312 | -37.3616
27.1212 | -35.6315 | -37.3669
27.1273 | -35.6318 | -37.3731
27.1334 | -35.6319 | -37.378
27.1395 | -35.6323 | -37.3828
27.1456 | -35.6327 | -37.3878
27.1518 | -35.633 | -37.3935
27.1579 | -35.633 | -37.3984
27.164 | -35.6334 | -37.4042
27.1701 | -35.6336 | -37.4087
27.1762 | -35.6338 | -37.4137
27.1823 | -35.6341 | -37.4194
27.1884 | -35.6345 | -37.4254
27.1945 | -35.6349 | -37.4313
27.2006 | -35.6351 | -37.4363
27.2067 | -35.6355 | -37.4418
27.2128 | -35.6358 | -37.4475
27.2189 | -35.636 | -37.4528
27.225 | -35.6364 | -37.459
27.2311 | -35.6367 | -37.4644
27.2372 | -35.6371 | -37.4694
27.2433 | -35.6373 | -37.4752
27.2494 | -35.6377 | -37.4795
27.2555 | -35.638 | -37.4846
27.2616 | -35.6382 | -37.4896
27.2677 | -35.6386 | -37.4947
27.2738 | -35.639 | -37.5006
27.2799 | -35.6394 | -37.5062
27.286 | -35.6397 | -37.5115
27.2921 | -35.64 | -37.5178
27.2983 | -35.6403 | -37.5229
27.3044 | -35.6406 | -37.528
27.3105 | -35.6408 | -37.5339
27.3166 | -35.6411 | -37.5393
27.3227 | -35.6415 | -37.5451
27.3288 | -35.6416 | -37.5505
27.3349 | -35.6419 | -37.5559
27.341 | -35.6422 | -37.5615
27.3471 | -35.6425 | -37.5667
27.3532 | -35.6428 | -37.5707
27.3593 | -35.6432 | -37.5762
27.3654 | -35.6436 | -37.5813
27.3715 | -35.644 | -37.5869
27.3776 | -35.6442 | -37.5921
27.3837 | -35.6444 | -37.5975
27.3898 | -35.6448 | -37.6032
27.3959 | -35.6452 | -37.6096
27.402 | -35.6454 | -37.614
27.4081 | -35.6458 | -37.6181
27.4142 | -35.6461 | -37.6234
27.4203 | -35.6465 | -37.6291
27.4264 | -35.6468 | -37.6348
27.4325 | -35.6471 | -37.6402
27.4387 | -35.6475 | -37.6452
27.4448 | -35.6478 | -37.6501
27.4509 | -35.6481 | -37.6559
27.457 | -35.6486 | -37.6611
27.4631 | -35.6489 | -37.6662
27.4692 | -35.6492 | -37.6715
27.4753 | -35.6495 | -37.6764
27.4814 | -35.6497 | -37.6818
27.4875 | -35.65 | -37.6875
27.4936 | -35.6503 | -37.6931
27.4997 | -35.6507 | -37.6989
27.5058 | -35.651 | -37.7036
27.5119 | -35.6513 | -37.7085
27.518 | -35.6514 | -37.7135
27.5241 | -35.6519 | -37.7187
27.5302 | -35.6523 | -37.7238
27.5363 | -35.6526 | -37.7289
27.5424 | -35.653 | -37.735
27.5485 | -35.6533 | -37.7406
27.5546 | -35.6538 | -37.7457
27.5607 | -35.6541 | -37.7509
27.5668 | -35.6545 | -37.7555
27.5729 | -35.6548 | -37.76
27.5791 | -35.6552 | -37.7644
27.5852 | -35.6555 | -37.7697
27.5913 | -35.656 | -37.775
27.5974 | -35.6562 | -37.78
27.6035 | -35.6565 | -37.7852
27.6096 | -35.6567 | -37.7899
27.6157 | -35.6571 | -37.7956
27.6218 | -35.6574 | -37.8015
27.6279 | -35.6577 | -37.807
27.634 | -35.6581 | -37.8121
27.6401 | -35.6583 | -37.8171
27.6462 | -35.6587 | -37.8223
27.6523 | -35.6589 | -37.8277
27.6584 | -35.6591 | -37.8331
27.6645 | -35.6595 | -37.838
27.6706 | -35.6599 | -37.8428
27.6767 | -35.6602 | -37.8486
27.6828 | -35.6605 | -37.8539
27.6889 | -35.6609 | -37.8598
27.695 | -35.6611 | -37.8642
27.7011 | -35.6614 | -37.8701
27.7072 | -35.6616 | -37.8762
27.7133 | -35.6619 | -37.8826
27.7194 | -35.6622 | -37.8883
27.7256 | -35.6625 | -37.8935
27.7317 | -35.6628 | -37.898
27.7378 | -35.6632 | -37.9043
27.7439 | -35.6635 | -37.9104
27.75 | -35.6639 | -37.9161
27.7561 | -35.6641 | -37.9223
27.7622 | -35.6643 | -37.9269
27.7683 | -35.6646 | -37.9333
27.7744 | -35.6648 | -37.9391
27.7805 | -35.6651 | -37.9446
27.7866 | -35.6655 | -37.9504
27.7927 | -35.6658 | -37.9558
27.7988 | -35.6662 | -37.9607
27.8049 | -35.6664 | -37.9657
27.811 | -35.6667 | -37.9713
27.8171 | -35.667 | -37.9765
27.8232 | -35.6673 | -37.9827
27.8293 | -35.6676 | -37.9879
27.8354 | -35.668 | -37.9929
27.8415 | -35.6683 | -37.9988
27.8476 | -35.6686 | -38.0039
27.8537 | -35.6689 | -38.0088
27.8598 | -35.6691 | -38.0134
27.866 | -35.6695 | -38.0193
27.8721 | -35.6698 | -38.0245
27.8782 | -35.6701 | -38.0299
27.8843 | -35.6703 | -38.0357
27.8904 | -35.6707 | -38.041
27.8965 | -35.6709 | -38.0467
27.9026 | -35.6711 | -38.0523
27.9087 | -35.6715 | -38.0576
27.9148 | -35.6718 | -38.0625
27.9209 | -35.672 | -38.0681
27.927 | -35.6723 | -38.0728
27.9331 | -35.6725 | -38.0786
27.9392 | -35.6728 | -38.084
27.9453 | -35.6731 | -38.0889
27.9514 | -35.6733 | -38.0942
27.9575 | -35.6736 | -38.099
27.9636 | -35.6738 | -38.1042
27.9697 | -35.674 | -38.1093
27.9758 | -35.6743 | -38.1149
27.9819 | -35.6747 | -38.1209
27.988 | -35.6749 | -38.1259
27.9941 | -35.6752 | -38.1309
28.0002 | -35.6755 | -38.1363
28.0063 | -35.6757 | -38.142
28.0125 | -35.676 | -38.147
28.0186 | -35.6762 | -38.1522
28.0247 | -35.6766 | -38.1573
28.0308 | -35.6768 | -38.163
28.0369 | -35.6772 | -38.1679
28.043 | -35.6775 | -38.1732
28.0491 | -35.6778 | -38.1783
28.0552 | -35.678 | -38.1833
28.0613 | -35.6784 | -38.1889
28.0674 | -35.6787 | -38.1943
28.0735 | -35.6789 | -38.1997

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.3921 | -42.073
21.4908 | -35.4876 | -35.1041
16.4948 | -34.8649 | -27.7063
12.6602 | -34.4972 | -21.4346
9.7171 | -34.2809 | -16.5195
7.4582 | -34.1318 | -12.5408
5.7244 | -34.0251 | -9.3182
4.3936 | -33.9601 | -6.7102
3.3722 | -33.9452 | -4.66
2.5883 | -33.9373 | -3.1089
1.9866 | -33.935 | -1.867
1.5248 | -33.9366 | -0.87091
1.1703 | -33.9424 | -0.21936
1 | -33.9387 | 0.10308
0.89822 | -33.9308 | 0.37322
0.68941 | -33.9192 | 0.95226
0.52919 | -33.9284 | 1.402
0.40612 | -33.9248 | 2.0088
0.31171 | -33.9383 | 2.4582
0.23928 | -33.9332 | 3.1809
0.18358 | -33.9548 | 3.5917
0.14091 | -33.9577 | 4.4712
0.10819 | -33.9816 | 5.6234
0.083017 | -34.0063 | 7.0928
0.063705 | -34.0435 | 8.9646
0.048923 | -34.1164 | 11.4495
0.037479 | -34.2172 | 14.6365
0.028801 | -34.3981 | 18.6567
0.022078 | -34.7079 | 24.0834
0.017023 | -35.2052 | 30.5354
0.013018 | -36.0169 | 38.2425
0.010014 | -37.1382 | 46.1689

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003711 | -6.3801
0.007423 | -25.3028
0.011134 | -34.7967
0.014846 | -40.0814
0.018557 | -41.9373
0.022269 | -43.7639
0.02598 | -46.5829
0.029691 | -49.9123
0.033403 | -49.5177
0.037114 | -50.5193
0.040826 | -51.2253
0.044537 | -51.5874
0.048249 | -52.9673
0.05196 | -54.9369
0.055672 | -56.3701
0.059383 | -57.2147
0.063094 | -57.4186
0.066806 | -57.7865
0.070517 | -58.1779
0.074229 | -59.4801
0.07794 | -61.6395
0.081652 | -61.8944
0.085363 | -62.8523
0.089074 | -62.3125
0.092786 | -63.4258
0.096497 | -63.715
0.10021 | -63.7769
0.10392 | -64.4231
0.10763 | -64.2751
0.11134 | -63.7073
0.11506 | -63.7587
0.11877 | -63.7765
0.12248 | -64.8811
0.12619 | -67.0996
0.1299 | -66.7683
0.13361 | -66.293
0.13732 | -67.1434
0.14103 | -67.9014
0.14475 | -67.2769
0.14846 | -67.8768
0.15217 | -68.6291
0.15588 | -69.5962
0.15959 | -70.2183
0.1633 | -69.0504
0.16701 | -68.5678
0.17073 | -69.2649
0.17444 | -69.668
0.17815 | -69.6755
0.18186 | -69.8824
0.18557 | -70.5551
0.18928 | -69.458
0.19299 | -70.5391
0.19671 | -72.1783
0.20042 | -72.3107
0.20413 | -71.6004
0.20784 | -71.1863
0.21155 | -71.043
0.21526 | -71.1617
0.21898 | -70.8294
0.22269 | -72.1574
0.2264 | -72.6269
0.23011 | -72.2946
0.23382 | -73.0231
0.23753 | -72.4415
0.24124 | -71.8125
0.24496 | -73.3641
0.24867 | -74.4611
0.25238 | -74.9378
0.25609 | -74.7785
0.2598 | -75.0106
0.26351 | -74.656
0.26722 | -74.7363
0.27093 | -75.2438
0.27465 | -73.8963
0.27836 | -72.5081
0.28207 | -74.0996
0.28578 | -74.6287
0.28949 | -74.185
0.2932 | -74.7671
0.29691 | -75.1734
0.30063 | -75.0928
0.30434 | -76.4509
0.30805 | -76.7082
0.31176 | -76.3363
0.31547 | -75.547
0.31918 | -75.7963
0.32289 | -76.0944
0.32661 | -75.9837
0.33032 | -76.4303
0.33403 | -75.9504
0.33774 | -76.4047
0.34145 | -77.2461
0.34516 | -77.1973
0.34887 | -77.1981
0.35259 | -76.7418
0.3563 | -76.4563
0.36001 | -76.8547
0.36372 | -76.5718
0.36743 | -77.7467
0.37114 | -78.2325
0.37485 | -78.0823
0.37857 | -78.1338
0.38228 | -78.1922
0.38599 | -77.0124
0.3897 | -77.5404
0.39341 | -78.5959
0.39712 | -78.8428
0.40083 | -78.439
0.40455 | -77.9746
0.40826 | -78.6335
0.41197 | -79.0926
0.41568 | -78.4418
0.41939 | -77.9118
0.4231 | -77.8572
0.42681 | -78.1726
0.43053 | -78.5677
0.43424 | -78.2256
0.43795 | -78.0334
0.44166 | -78.7054
0.44537 | -79.3262
0.44908 | -79.4467
0.4528 | -79.1196
0.45651 | -79.4213
0.46022 | -80.0462
0.46393 | -79.5458
0.46764 | -79.5707
0.47135 | -79.875
0.47506 | -79.1938
0.47878 | -79.6434
0.48249 | -79.7026
0.4862 | -80.032
0.48991 | -79.8834
0.49362 | -80.1357
0.49733 | -79.025
0.50104 | -79.5707
0.50475 | -80.8432
0.50847 | -80.6145
0.51218 | -80.6479
0.51589 | -81.1669
0.5196 | -81.7462
0.52331 | -81.3076
0.52702 | -80.5892
0.53073 | -79.3712
0.53445 | -79.81
0.53816 | -80.325
0.54187 | -80.846
0.54558 | -81.152
0.54929 | -80.8943
0.553 | -80.9414
0.55671 | -81.5529
0.56043 | -82.0416
0.56414 | -82.2635
0.56785 | -82.557
0.57156 | -82.6353
0.57527 | -82.2914
0.57898 | -82.25
0.58269 | -82.7406
0.58641 | -83.0088
0.59012 | -82.8793
0.59383 | -82.8669
0.59754 | -82.909
0.60125 | -82.565
0.60496 | -82.0476
0.60867 | -82.0348
0.61239 | -82.3235
0.6161 | -82.3011
0.61981 | -82.1164
0.62352 | -82.2744
0.62723 | -82.5529
0.63094 | -82.6833
0.63466 | -83.0974
0.63837 | -83.7299
0.64208 | -84.1399
0.64579 | -84.0311
0.6495 | -83.4431
0.65321 | -82.9107
0.65692 | -82.8984
0.66064 | -83.3986
0.66435 | -83.7872
0.66806 | -83.5472
0.67177 | -83.3058
0.67548 | -83.3519
0.67919 | -83.4407
0.6829 | -83.1446
0.68662 | -83.1836
0.69033 | -83.9253
0.69404 | -84.9403
0.69775 | -85.1286
0.70146 | -84.7176
0.70517 | -84.5888
0.70888 | -84.4652
0.7126 | -84.4164
0.71631 | -84.2992
0.72002 | -84.2927
0.72373 | -84.5002
0.72744 | -84.475
0.73115 | -84.614
0.73486 | -85.0579
0.73858 | -85.3358
0.74229 | -85.4778
0.746 | -85.6284
0.74971 | -85.3384
0.75342 | -84.8585
0.75713 | -84.3702
0.76084 | -83.9898
0.76456 | -83.838
0.76827 | -83.8702
0.77198 | -84.2409
0.77569 | -84.4818
0.7794 | -84.5029
0.78311 | -84.774
0.78682 | -85.2562
0.79054 | -85.3699
0.79425 | -84.9625
0.79796 | -84.9336
0.80167 | -85.236
0.80538 | -85.1506
0.80909 | -84.8351
0.8128 | -84.6526
0.81652 | -84.9761
0.82023 | -85.6699
0.82394 | -85.8526
0.82765 | -85.3444
0.83136 | -84.9978
0.83507 | -85.2437
0.83879 | -85.8775
0.8425 | -86.3054
0.84621 | -86.1934
0.84992 | -85.9774
0.85363 | -85.9733
0.85734 | -86.0539
0.86105 | -85.9969
0.86477 | -85.561
0.86848 | -85.23
0.87219 | -85.4226
0.8759 | -85.7185
0.87961 | -86.2072
0.88332 | -86.372
0.88703 | -85.7554
0.89075 | -85.3225
0.89446 | -85.8123
0.89817 | -86.4116
0.90188 | -86.5187
0.90559 | -86.3566
0.9093 | -85.9575
0.91301 | -85.6572
0.91673 | -85.8881
0.92044 | -86.2732
0.92415 | -86.183
0.92786 | -85.8973
0.93157 | -86.1812
0.93528 | -86.6021
0.93899 | -86.6493
0.94271 | -86.727
0.94642 | -86.9465
0.95013 | -87.1903
0.95384 | -87.1386
0.95755 | -86.925
0.96126 | -86.6793
0.96497 | -86.5855
0.96869 | -86.7119
0.9724 | -86.7988
0.97611 | -86.7857
0.97982 | -86.8926
0.98353 | -87.0596
0.98724 | -87.2151
0.99095 | -87.0969
0.99467 | -86.6735
0.99838 | -86.6975
1.0021 | -86.9436
1.0058 | -87.2646
1.0095 | -87.2605
1.0132 | -86.8791
1.0169 | -86.8248
1.0206 | -86.8295
1.0244 | -86.9912
1.0281 | -87.2077
1.0318 | -87.5074
1.0355 | -88.0668
1.0392 | -88.1999
1.0429 | -88.0903
1.0466 | -88.1478
1.0503 | -88.4044
1.054 | -88.7544
1.0578 | -88.8988
1.0615 | -88.6568
1.0652 | -87.8872
1.0689 | -87.8529
1.0726 | -88.0542
1.0763 | -88.4565
1.08 | -88.7676
1.0837 | -88.7317
1.0875 | -88.5113
1.0912 | -88.4002
1.0949 | -88.2695
1.0986 | -87.981
1.1023 | -87.7921
1.106 | -87.9032
1.1097 | -88.1089
1.1134 | -88.162
1.1171 | -88.2074
1.1209 | -88.3913
1.1246 | -88.5574
1.1283 | -88.5494
1.132 | -88.5
1.1357 | -88.5705
1.1394 | -88.6648
1.1431 | -88.5664
1.1468 | -88.3311
1.1505 | -88.271
1.1543 | -88.4816
1.158 | -88.7507
1.1617 | -88.9677
1.1654 | -89.058
1.1691 | -88.9602
1.1728 | -88.898
1.1765 | -88.931
1.1802 | -88.7666
1.1839 | -88.4263
1.1877 | -88.2435
1.1914 | -88.4271
1.1951 | -88.8782
1.1988 | -89.2531
1.2025 | -89.3382
1.2062 | -89.2261
1.2099 | -89.06
1.2136 | -89.0169
1.2174 | -89.2217
1.2211 | -89.5237
1.2248 | -89.6129
1.2285 | -89.533
1.2322 | -89.5126
1.2359 | -89.5409
1.2396 | -89.5274
1.2433 | -89.5541
1.247 | -89.6446
1.2508 | -89.5823
1.2545 | -89.3067
1.2582 | -89.1789
1.2619 | -89.4434
1.2656 | -89.8629
1.2693 | -90.0924
1.273 | -90.1244
1.2767 | -90.1094
1.2804 | -90.1625
1.2842 | -90.2155
1.2879 | -90.1574
1.2916 | -90.1019
1.2953 | -90.0904
1.299 | -90.0103
1.3027 | -89.8139
1.3064 | -89.5943
1.3101 | -89.5227
1.3138 | -89.6341
1.3176 | -89.7447
1.3213 | -89.7178
1.325 | -89.6218
1.3287 | -89.5209
1.3324 | -89.4253
1.3361 | -89.3855
1.3398 | -89.3941
1.3435 | -89.3921
1.3473 | -89.3698
1.351 | -89.3657
1.3547 | -89.3607
1.3584 | -89.3914
1.3621 | -89.63
1.3658 | -90.0505
1.3695 | -90.3223
1.3732 | -90.2364
1.3769 | -90.1374
1.3807 | -90.3568
1.3844 | -90.5834
1.3881 | -90.4508
1.3918 | -90.1964
1.3955 | -90.127
1.3992 | -90.2776
1.4029 | -90.4997
1.4066 | -90.614
1.4103 | -90.5675
1.4141 | -90.4529
1.4178 | -90.3963
1.4215 | -90.4349
1.4252 | -90.499
1.4289 | -90.5244
1.4326 | -90.5231
1.4363 | -90.4877
1.44 | -90.4135
1.4437 | -90.3845
1.4475 | -90.4278
1.4512 | -90.4644
1.4549 | -90.5405
1.4586 | -90.722
1.4623 | -90.9016
1.466 | -91.0412
1.4697 | -91.1688
1.4734 | -91.2657
1.4772 | -91.3362
1.4809 | -91.3314
1.4846 | -91.2908
1.4883 | -91.3602
1.492 | -91.4839
1.4957 | -91.4414
1.4994 | -91.1596
1.5031 | -90.767
1.5068 | -90.4703
1.5106 | -90.3877
1.5143 | -90.4112
1.518 | -90.422
1.5217 | -90.6047
1.5254 | -91.0403
1.5291 | -91.3201
1.5328 | -91.1069
1.5365 | -90.7924
1.5402 | -90.8237
1.544 | -91.1557
1.5477 | -91.439
1.5514 | -91.5011
1.5551 | -91.4479
1.5588 | -91.4115
1.5625 | -91.4025
1.5662 | -91.3156
1.5699 | -91.1707
1.5736 | -91.1995
1.5774 | -91.4518
1.5811 | -91.649
1.5848 | -91.6633
1.5885 | -91.6878
1.5922 | -91.7673
1.5959 | -91.7206
1.5996 | -91.4982
1.6033 | -91.2709
1.6071 | -91.2103
1.6108 | -91.3401
1.6145 | -91.5659
1.6182 | -91.7321
1.6219 | -91.7405
1.6256 | -91.6476
1.6293 | -91.514
1.633 | -91.4182
1.6367 | -91.3926
1.6405 | -91.3953
1.6442 | -91.3908
1.6479 | -91.4371
1.6516 | -91.571
1.6553 | -91.7475
1.659 | -91.8092
1.6627 | -91.7272
1.6664 | -91.6115
1.6701 | -91.5279
1.6739 | -91.4634
1.6776 | -91.3791
1.6813 | -91.3494
1.685 | -91.4991
1.6887 | -91.8206
1.6924 | -92.2067
1.6961 | -92.5354
1.6998 | -92.7472
1.7035 | -92.822
1.7073 | -92.7408
1.711 | -92.5418
1.7147 | -92.3059
1.7184 | -92.1474
1.7221 | -92.1359
1.7258 | -92.2577
1.7295 | -92.4544
1.7332 | -92.6514
1.737 | -92.7409
1.7407 | -92.6696
1.7444 | -92.4985
1.7481 | -92.3565
1.7518 | -92.3143
1.7555 | -92.3218
1.7592 | -92.3184
1.7629 | -92.3438
1.7666 | -92.4704
1.7704 | -92.6905
1.7741 | -92.8935
1.7778 | -92.9605
1.7815 | -92.8285
1.7852 | -92.5456
1.7889 | -92.2093
1.7926 | -91.8883
1.7963 | -91.6536
1.8 | -91.5622
1.8038 | -91.6094
1.8075 | -91.7407
1.8112 | -91.9373
1.8149 | -92.2125
1.8186 | -92.5078
1.8223 | -92.7281
1.826 | -92.8605
1.8297 | -92.9715
1.8334 | -93.0367
1.8372 | -92.9404
1.8409 | -92.6793
1.8446 | -92.4459
1.8483 | -92.4366
1.852 | -92.6886
1.8557 | -93.0761
1.8594 | -93.3657
1.8631 | -93.3698
1.8669 | -93.1457
1.8706 | -92.9246
1.8743 | -92.8538
1.878 | -92.9253
1.8817 | -93.0694
1.8854 | -93.2408
1.8891 | -93.3648
1.8928 | -93.3771
1.8965 | -93.2906
1.9003 | -93.1585
1.904 | -93.0119
1.9077 | -92.8636
1.9114 | -92.7401
1.9151 | -92.7026
1.9188 | -92.7814
1.9225 | -92.9528
1.9262 | -93.1472
1.9299 | -93.2806
1.9337 | -93.3485
1.9374 | -93.4379
1.9411 | -93.6228
1.9448 | -93.8419
1.9485 | -93.9398
1.9522 | -93.8139
1.9559 | -93.5394
1.9596 | -93.2708
1.9633 | -93.1522
1.9671 | -93.2327
1.9708 | -93.4679
1.9745 | -93.7885
1.9782 | -94.0717
1.9819 | -94.1562
1.9856 | -94.0006
1.9893 | -93.7192
1.993 | -93.5133
1.9968 | -93.4871
2.0005 | -93.5944
2.0042 | -93.7333
2.0079 | -93.826
2.0116 | -93.8824
2.0153 | -93.9181
2.019 | -93.9377
2.0227 | -93.9626
2.0264 | -94.0075
2.0302 | -94.0599
2.0339 | -94.0541
2.0376 | -93.9128
2.0413 | -93.6467
2.045 | -93.3727
2.0487 | -93.2081
2.0524 | -93.1773
2.0561 | -93.2641
2.0598 | -93.4585
2.0636 | -93.7208
2.0673 | -93.9618
2.071 | -94.0795
2.0747 | -94.0879
2.0784 | -94.0349
2.0821 | -93.9437
2.0858 | -93.8079
2.0895 | -93.6167
2.0932 | -93.4308
2.097 | -93.3169
2.1007 | -93.2908
2.1044 | -93.3382
2.1081 | -93.4511
2.1118 | -93.6168
2.1155 | -93.759
2.1192 | -93.7934
2.1229 | -93.7516
2.1267 | -93.7106
2.1304 | -93.7286
2.1341 | -93.8417
2.1378 | -94.0672
2.1415 | -94.2907
2.1452 | -94.4467
2.1489 | -94.4785
2.1526 | -94.3689
2.1563 | -94.175
2.1601 | -93.9512
2.1638 | -93.7487
2.1675 | -93.6155
2.1712 | -93.5566
2.1749 | -93.5504
2.1786 | -93.5696
2.1823 | -93.5901
2.186 | -93.6177
2.1897 | -93.6639
2.1935 | -93.718
2.1972 | -93.7909
2.2009 | -93.8871
2.2046 | -93.9868
2.2083 | -94.0405
2.212 | -94.0234
2.2157 | -93.9311
2.2194 | -93.7912
2.2231 | -93.6599
2.2269 | -93.5741
2.2306 | -93.5617
2.2343 | -93.6188
2.238 | -93.7137
2.2417 | -93.8173
2.2454 | -93.9021
2.2491 | -93.9465
2.2528 | -93.955
2.2566 | -93.9511
2.2603 | -93.9338
2.264 | -93.8962
2.2677 | -93.8654
2.2714 | -93.869
2.2751 | -93.9486
2.2788 | -94.0887
2.2825 | -94.2316
2.2862 | -94.3312
2.29 | -94.4244
2.2937 | -94.5847
2.2974 | -94.8529
2.3011 | -95.2047
2.3048 | -95.5553
2.3085 | -95.8033
2.3122 | -95.8758
2.3159 | -95.7368
2.3196 | -95.4672
2.3234 | -95.1502
2.3271 | -94.849
2.3308 | -94.624
2.3345 | -94.5321
2.3382 | -94.5932
2.3419 | -94.7509
2.3456 | -94.889
2.3493 | -94.9256
2.3531 | -94.8385
2.3568 | -94.686
2.3605 | -94.5439
2.3642 | -94.4707
2.3679 | -94.4656
2.3716 | -94.496
2.3753 | -94.5138
2.379 | -94.5113
2.3827 | -94.5286
2.3865 | -94.6055
2.3902 | -94.7489
2.3939 | -94.9361
2.3976 | -95.0863
2.4013 | -95.1306
2.405 | -95.0679
2.4087 | -94.9616
2.4124 | -94.8885
2.4161 | -94.9051
2.4199 | -94.9853
2.4236 | -95.0434
2.4273 | -95.028
2.431 | -94.9523
2.4347 | -94.8732
2.4384 | -94.849
2.4421 | -94.8985
2.4458 | -95.0061
2.4495 | -95.1374
2.4533 | -95.2451
2.457 | -95.3156
2.4607 | -95.3455
2.4644 | -95.3553
2.4681 | -95.3549
2.4718 | -95.3513
2.4755 | -95.3351
2.4792 | -95.3025
2.483 | -95.2905
2.4867 | -95.3455
2.4904 | -95.4656
2.4941 | -95.6104
2.4978 | -95.6776
2.5015 | -95.6195
2.5052 | -95.4687
2.5089 | -95.3178
2.5126 | -95.262
2.5164 | -95.3228
2.5201 | -95.4292
2.5238 | -95.475
2.5275 | -95.3741
2.5312 | -95.1454
2.5349 | -94.8821
2.5386 | -94.6938
2.5423 | -94.6139
2.546 | -94.637
2.5498 | -94.7372
2.5535 | -94.864
2.5572 | -94.9941
2.5609 | -95.1163
2.5646 | -95.1983
2.5683 | -95.2477
2.572 | -95.2937
2.5757 | -95.3538
2.5794 | -95.4302
2.5832 | -95.5067
2.5869 | -95.5427
2.5906 | -95.5499
2.5943 | -95.5579
2.598 | -95.5728
2.6017 | -95.6263
2.6054 | -95.7379
2.6091 | -95.8756
2.6129 | -96.0001
2.6166 | -96.1199
2.6203 | -96.2536
2.624 | -96.4095
2.6277 | -96.563
2.6314 | -96.6366
2.6351 | -96.5887
2.6388 | -96.4058
2.6425 | -96.1001
2.6463 | -95.7472
2.65 | -95.4241
2.6537 | -95.1988
2.6574 | -95.0988
2.6611 | -95.1438
2.6648 | -95.3027
2.6685 | -95.52
2.6722 | -95.7527
2.6759 | -95.9503
2.6797 | -96.1387
2.6834 | -96.2739
2.6871 | -96.3317
2.6908 | -96.3028
2.6945 | -96.2225
2.6982 | -96.1152
2.7019 | -96.0058
2.7056 | -95.8902
2.7093 | -95.7798
2.7131 | -95.6905
2.7168 | -95.6308
2.7205 | -95.606
2.7242 | -95.6117
2.7279 | -95.6235
2.7316 | -95.633
2.7353 | -95.6229
2.739 | -95.5978
2.7428 | -95.5632
2.7465 | -95.5502
2.7502 | -95.5711
2.7539 | -95.6205
2.7576 | -95.7076
2.7613 | -95.8277
2.765 | -95.9882
2.7687 | -96.1664
2.7724 | -96.3338
2.7762 | -96.4498
2.7799 | -96.4808
2.7836 | -96.4359
2.7873 | -96.3319
2.791 | -96.2067
2.7947 | -96.0962
2.7984 | -96.0366
2.8021 | -96.0391
2.8058 | -96.0866
2.8096 | -96.1384
2.8133 | -96.1642
2.817 | -96.1455
2.8207 | -96.0801
2.8244 | -95.968
2.8281 | -95.8281
2.8318 | -95.7022
2.8355 | -95.653
2.8392 | -95.7008
2.843 | -95.845
2.8467 | -96.0417
2.8504 | -96.2324
2.8541 | -96.3695
2.8578 | -96.444
2.8615 | -96.4528
2.8652 | -96.409
2.8689 | -96.3253
2.8727 | -96.225
2.8764 | -96.124
2.8801 | -96.0584
2.8838 | -96.0274
2.8875 | -96.0456
2.8912 | -96.1015
2.8949 | -96.1744
2.8986 | -96.239
2.9023 | -96.2829
2.9061 | -96.2991
2.9098 | -96.2838
2.9135 | -96.2782
2.9172 | -96.2903
2.9209 | -96.3303
2.9246 | -96.4032
2.9283 | -96.5067
2.932 | -96.6289
2.9357 | -96.7203
2.9395 | -96.728
2.9432 | -96.65
2.9469 | -96.5233
2.9506 | -96.3936
2.9543 | -96.3199
2.958 | -96.3148
2.9617 | -96.3668
2.9654 | -96.4653
2.9691 | -96.5781
2.9729 | -96.6769
2.9766 | -96.7241
2.9803 | -96.7254
2.984 | -96.6857
2.9877 | -96.6369
2.9914 | -96.5856
2.9951 | -96.5396
2.9988 | -96.5012
3.0026 | -96.4611
3.0063 | -96.4193
3.01 | -96.3724
3.0137 | -96.3445
3.0174 | -96.3556
3.0211 | -96.4088
3.0248 | -96.5236
3.0285 | -96.693
3.0322 | -96.895
3.036 | -97.09
3.0397 | -97.2575
3.0434 | -97.3578
3.0471 | -97.3741
3.0508 | -97.3234
3.0545 | -97.244
3.0582 | -97.1559
3.0619 | -97.0675
3.0656 | -96.9799
3.0694 | -96.8894
3.0731 | -96.7853
3.0768 | -96.6819
3.0805 | -96.586
3.0842 | -96.5174
3.0879 | -96.4939
3.0916 | -96.5097
3.0953 | -96.5449
3.099 | -96.5736
3.1028 | -96.5797
3.1065 | -96.5761
3.1102 | -96.5902
3.1139 | -96.6325
3.1176 | -96.7056
3.1213 | -96.7806
3.125 | -96.8282
3.1287 | -96.8369
3.1325 | -96.7821
3.1362 | -96.6722
3.1399 | -96.5545
3.1436 | -96.4544
3.1473 | -96.3867
3.151 | -96.3624
3.1547 | -96.3755
3.1584 | -96.4086
3.1621 | -96.4471
3.1659 | -96.4871
3.1696 | -96.5432
3.1733 | -96.6291
3.177 | -96.7563
3.1807 | -96.9119
3.1844 | -97.0451
3.1881 | -97.1021
3.1918 | -97.0726
3.1955 | -96.977
3.1993 | -96.849
3.203 | -96.7131
3.2067 | -96.6102
3.2104 | -96.6151
3.2141 | -96.5964
3.2178 | -96.6114
3.2215 | -96.6491
3.2252 | -96.6891
3.2289 | -96.7195
3.2327 | -96.7499
3.2364 | -96.7793
3.2401 | -96.8046
3.2438 | -96.8242
3.2475 | -96.8474
3.2512 | -96.8807
3.2549 | -96.944
3.2586 | -97.0252
3.2624 | -97.1147
3.2661 | -97.2091
3.2698 | -97.2744
3.2735 | -97.2893
3.2772 | -97.2666
3.2809 | -97.2225
3.2846 | -97.1826
3.2883 | -97.1682
3.292 | -97.1908
3.2958 | -97.248
3.2995 | -97.3335
3.3032 | -97.4222
3.3069 | -97.4963
3.3106 | -97.5335
3.3143 | -97.5293
3.318 | -97.4751
3.3217 | -97.3588
3.3254 | -97.2176
3.3292 | -97.0691
3.3329 | -96.9351
3.3366 | -96.8434
3.3403 | -96.7945
3.344 | -96.7902
3.3477 | -96.8368
3.3514 | -96.9238
3.3551 | -97.0248
3.3588 | -97.1258
3.3626 | -97.212
3.3663 | -97.2748
3.37 | -97.322
3.3737 | -97.366
3.3774 | -97.4252
3.3811 | -97.4949
3.3848 | -97.5767
3.3885 | -97.6704
3.3923 | -97.7688
3.396 | -97.8618
3.3997 | -97.9569
3.4034 | -98.0379
3.4071 | -98.0923
3.4108 | -98.1345
3.4145 | -98.1573
3.4182 | -98.1682
3.4219 | -98.1819
3.4257 | -98.1892
3.4294 | -98.1868
3.4331 | -98.1596
3.4368 | -98.0999
3.4405 | -98.0155
3.4442 | -97.9163
3.4479 | -97.8285
3.4516 | -97.7762
3.4553 | -97.7577
3.4591 | -97.7936
3.4628 | -97.8711
3.4665 | -97.9681
3.4702 | -98.0748
3.4739 | -98.154
3.4776 | -98.1928
3.4813 | -98.1959
3.485 | -98.1809
3.4888 | -98.1603
3.4925 | -98.1449
3.4962 | -98.137
3.4999 | -98.1228
3.5036 | -98.0852
3.5073 | -98.0252
3.511 | -97.9514
3.5147 | -97.8771
3.5184 | -97.8309
3.5222 | -97.8153
3.5259 | -97.8182
3.5296 | -97.8472
3.5333 | -97.8865
3.537 | -97.9217
3.5407 | -97.9348
3.5444 | -97.9365
3.5481 | -97.927
3.5518 | -97.9097
3.5556 | -97.9052
3.5593 | -97.92
3.563 | -97.9568
3.5667 | -98.0106
3.5704 | -98.069
3.5741 | -98.1104
3.5778 | -98.1209
3.5815 | -98.0838
3.5852 | -97.9907
3.589 | -97.8721
3.5927 | -97.7654
3.5964 | -97.7067
3.6001 | -97.7119
3.6038 | -97.7869
3.6075 | -97.9148
3.6112 | -98.0767
3.6149 | -98.2433
3.6187 | -98.3818
3.6224 | -98.4708
3.6261 | -98.4967
3.6298 | -98.4733
3.6335 | -98.4223
3.6372 | -98.3786
3.6409 | -98.3645
3.6446 | -98.3923
3.6483 | -98.4547
3.6521 | -98.5249
3.6558 | -98.5757
3.6595 | -98.5961
3.6632 | -98.5884
3.6669 | -98.5571
3.6706 | -98.5434
3.6743 | -98.5671
3.678 | -98.6229
3.6817 | -98.7029
3.6855 | -98.7853
3.6892 | -98.8482
3.6929 | -98.8688
3.6966 | -98.8437
3.7003 | -98.7739
3.704 | -98.6861
3.7077 | -98.6001
3.7114 | -98.5403
3.7151 | -98.5185
3.7189 | -98.5315
3.7226 | -98.5796
3.7263 | -98.6371
3.73 | -98.6968
3.7337 | -98.7581
3.7374 | -98.8169
3.7411 | -98.8847
3.7448 | -98.9805
3.7485 | -99.0579
3.7523 | -99.1273
3.756 | -99.1702
3.7597 | -99.1906
3.7634 | -99.1776
3.7671 | -99.1422
3.7708 | -99.0965
3.7745 | -99.0463
3.7782 | -99.0056
3.782 | -98.9634
3.7857 | -98.9249
3.7894 | -98.8885
3.7931 | -98.8425
3.7968 | -98.7872
3.8005 | -98.7272
3.8042 | -98.6697
3.8079 | -98.6147
3.8116 | -98.5744
3.8154 | -98.5425
3.8191 | -98.5193
3.8228 | -98.5077
3.8265 | -98.514
3.8302 | -98.5385
3.8339 | -98.5822
3.8376 | -98.6362
3.8413 | -98.6728
3.845 | -98.6879
3.8488 | -98.6789
3.8525 | -98.6463
3.8562 | -98.6088
3.8599 | -98.5768
3.8636 | -98.5557
3.8673 | -98.5488
3.871 | -98.5533
3.8747 | -98.5621
3.8784 | -98.5668
3.8822 | -98.563
3.8859 | -98.5405
3.8896 | -98.5175
3.8933 | -98.5014
3.897 | -98.4942
3.9007 | -98.511
3.9044 | -98.5616
3.9081 | -98.6297
3.9119 | -98.7047
3.9156 | -98.7834
3.9193 | -98.8492
3.923 | -98.9029
3.9267 | -98.9442
3.9304 | -98.9718
3.9341 | -98.9923
3.9378 | -98.999
3.9415 | -98.9964
3.9453 | -98.9892
3.949 | -98.9834
3.9527 | -98.9652
3.9564 | -98.9433
3.9601 | -98.9179
3.9638 | -98.8789
3.9675 | -98.8317
3.9712 | -98.787
3.9749 | -98.7517
3.9787 | -98.7416
3.9824 | -98.7691
3.9861 | -98.8326
3.9898 | -98.9322
3.9935 | -99.0522
3.9972 | -99.1785
4.0009 | -99.3012
4.0046 | -99.397
4.0084 | -99.4567
4.0121 | -99.4708
4.0158 | -99.4388
4.0195 | -99.3684
4.0232 | -99.2743
4.0269 | -99.1815
4.0306 | -99.0941
4.0343 | -99.0316
4.038 | -99.0098
4.0418 | -99.0233
4.0455 | -99.0678
4.0492 | -99.1342
4.0529 | -99.2096
4.0566 | -99.2865
4.0603 | -99.347
4.064 | -99.39
4.0677 | -99.4171
4.0714 | -99.4166
4.0752 | -99.4046
4.0789 | -99.3818
4.0826 | -99.3442
4.0863 | -99.2981
4.09 | -99.2453
4.0937 | -99.1916
4.0974 | -99.1418
4.1011 | -99.0981
4.1048 | -99.0756
4.1086 | -99.083
4.1123 | -99.1128
4.116 | -99.166
4.1197 | -99.2281
4.1234 | -99.2777
4.1271 | -99.2945
4.1308 | -99.2882
4.1345 | -99.2684
4.1383 | -99.238
4.142 | -99.2113
4.1457 | -99.1954
4.1494 | -99.1919
4.1531 | -99.1856
4.1568 | -99.1768
4.1605 | -99.1512
4.1642 | -99.1134
4.1679 | -99.0708
4.1717 | -99.0319
4.1754 | -99.016
4.1791 | -99.0207
4.1828 | -99.0464
4.1865 | -99.0925
4.1902 | -99.1557
4.1939 | -99.2222
4.1976 | -99.2982
4.2013 | -99.3786
4.2051 | -99.4539
4.2088 | -99.5274
4.2125 | -99.5999
4.2162 | -99.6867
4.2199 | -99.7892
4.2236 | -99.9014
4.2273 | -100.0091
4.231 | -100.0982
4.2347 | -100.1576
4.2385 | -100.1738
4.2422 | -100.1529
4.2459 | -100.1027
4.2496 | -100.0252
4.2533 | -99.9291
4.257 | -99.8168
4.2607 | -99.6965
4.2644 | -99.585
4.2682 | -99.489
4.2719 | -99.4126
4.2756 | -99.3644
4.2793 | -99.3513
4.283 | -99.4372
4.2867 | -99.5014
4.2904 | -99.5988
4.2941 | -99.7047
4.2978 | -99.8066
4.3016 | -99.8911
4.3053 | -99.9511
4.309 | -99.9808
4.3127 | -99.987
4.3164 | -99.9759
4.3201 | -99.9497
4.3238 | -99.9188
4.3275 | -99.8814
4.3312 | -99.8467
4.335 | -99.8329
4.3387 | -99.8256
4.3424 | -99.829
4.3461 | -99.8386
4.3498 | -99.8506
4.3535 | -99.8714
4.3572 | -99.8858
4.3609 | -99.9065
4.3646 | -99.9377
4.3684 | -99.9686
4.3721 | -99.9958
4.3758 | -100.0144
4.3795 | -100.019
4.3832 | -100.0016
4.3869 | -99.9544
4.3906 | -99.9013
4.3943 | -99.8384
4.3981 | -99.7851
4.4018 | -99.7319
4.4055 | -99.6859
4.4092 | -99.6654
4.4129 | -99.6627
4.4166 | -99.6757
4.4203 | -99.7011
4.424 | -99.7387
4.4277 | -99.7928
4.4315 | -99.8553
4.4352 | -99.9207
4.4389 | -99.9802
4.4426 | -100.0319
4.4463 | -100.0769
4.45 | -100.1189
4.4537 | -100.1598
4.4574 | -100.2109
4.4611 | -100.2582
4.4649 | -100.3121
4.4686 | -100.3546
4.4723 | -100.3988
4.476 | -100.4222
4.4797 | -100.4201
4.4834 | -100.4
4.4871 | -100.375
4.4908 | -100.3377
4.4945 | -100.2863
4.4983 | -100.2332
4.502 | -100.1789
4.5057 | -100.1295
4.5094 | -100.0877
4.5131 | -100.0534
4.5168 | -100.0233
4.5205 | -99.9986
4.5242 | -99.9859
4.528 | -99.9684
4.5317 | -99.9527
4.5354 | -99.9461
4.5391 | -99.9479
4.5428 | -99.9499
4.5465 | -99.9578
4.5502 | -99.9678
4.5539 | -99.9746
4.5576 | -99.9683
4.5614 | -99.963
4.5651 | -99.961
4.5688 | -99.9553
4.5725 | -99.9585
4.5762 | -99.9645
4.5799 | -99.9822
4.5836 | -100.0055
4.5873 | -100.0387
4.591 | -100.0815
4.5948 | -100.1233
4.5985 | -100.1685
4.6022 | -100.201
4.6059 | -100.2192
4.6096 | -100.2308
4.6133 | -100.2409
4.617 | -100.2439
4.6207 | -100.2355
4.6244 | -100.2224
4.6282 | -100.2132
4.6319 | -100.201
4.6356 | -100.1822
4.6393 | -100.1654
4.643 | -100.153
4.6467 | -100.1422
4.6504 | -100.1296
4.6541 | -100.1295
4.6579 | -100.1307
4.6616 | -100.1433
4.6653 | -100.1691
4.669 | -100.205
4.6727 | -100.2454
4.6764 | -100.2785
4.6801 | -100.3052
4.6838 | -100.3268
4.6875 | -100.3446
4.6913 | -100.3706
4.695 | -100.3983
4.6987 | -100.4236
4.7024 | -100.4573
4.7061 | -100.4919
4.7098 | -100.5246
4.7135 | -100.5456
4.7172 | -100.556
4.7209 | -100.5526
4.7247 | -100.5353
4.7284 | -100.5041
4.7321 | -100.4553
4.7358 | -100.4016
4.7395 | -100.3463
4.7432 | -100.293
4.7469 | -100.2475
4.7506 | -100.2069
4.7543 | -100.184
4.7581 | -100.1847
4.7618 | -100.2008
4.7655 | -100.2318
4.7692 | -100.2705
4.7729 | -100.3149
4.7766 | -100.3628
4.7803 | -100.3915
4.784 | -100.4069
4.7878 | -100.4114
4.7915 | -100.4055
4.7952 | -100.3829
4.7989 | -100.3578
4.8026 | -100.3253
4.8063 | -100.296
4.81 | -100.277
4.8137 | -100.2656
4.8174 | -100.2934
4.8212 | -100.3126
4.8249 | -100.3518
4.8286 | -100.401
4.8323 | -100.4635
4.836 | -100.5416
4.8397 | -100.626
4.8434 | -100.711
4.8471 | -100.7683
4.8508 | -100.8025
4.8546 | -100.8152
4.8583 | -100.7977
4.862 | -100.7603
4.8657 | -100.7098
4.8694 | -100.6554
4.8731 | -100.6053
4.8768 | -100.57
4.8805 | -100.5435
4.8842 | -100.5267
4.888 | -100.5203
4.8917 | -100.5139
4.8954 | -100.512
4.8991 | -100.5067
4.9028 | -100.4935
4.9065 | -100.4858
4.9102 | -100.4859
4.9139 | -100.4893
4.9177 | -100.4946
4.9214 | -100.5043
4.9251 | -100.523
4.9288 | -100.5402
4.9325 | -100.5597
4.9362 | -100.5783
4.9399 | -100.5841
4.9436 | -100.5791
4.9473 | -100.5721
4.9511 | -100.5585
4.9548 | -100.5384
4.9585 | -100.5168
4.9622 | -100.5039
4.9659 | -100.4975
4.9696 | -100.4957
4.9733 | -100.4983
4.977 | -100.505
4.9807 | -100.5084
4.9845 | -100.5119
4.9882 | -100.5133
4.9919 | -100.5202
4.9956 | -100.5392
4.9993 | -100.5627
5.003 | -100.5922
5.0067 | -100.6234
5.0104 | -100.6506
5.0141 | -100.6797
5.0179 | -100.7115
5.0216 | -100.7401
5.0253 | -100.7641
5.029 | -100.7816
5.0327 | -100.7931
5.0364 | -100.7999
5.0401 | -100.7961
5.0438 | -100.7905
5.0476 | -100.7926
5.0513 | -100.8012
5.055 | -100.8215
5.0587 | -100.8624
5.0624 | -100.9111
5.0661 | -100.9771
5.0698 | -101.0483
5.0735 | -101.1169
5.0772 | -101.1832
5.081 | -101.2476
5.0847 | -101.3036
5.0884 | -101.3554
5.0921 | -101.3984
5.0958 | -101.4392
5.0995 | -101.4795
5.1032 | -101.5051
5.1069 | -101.5217
5.1106 | -101.5328
5.1144 | -101.5396
5.1181 | -101.5454
5.1218 | -101.5476
5.1255 | -101.5478
5.1292 | -101.5529
5.1329 | -101.5628
5.1366 | -101.5773
5.1403 | -101.5945
5.1441 | -101.6038
5.1478 | -101.6171
5.1515 | -101.6223
5.1552 | -101.6204
5.1589 | -101.6181
5.1626 | -101.6192
5.1663 | -101.6183
5.17 | -101.6112
5.1737 | -101.6019
5.1775 | -101.5861
5.1812 | -101.5624
5.1849 | -101.5277
5.1886 | -101.4763
5.1923 | -101.4102
5.196 | -101.3437
5.1997 | -101.2861
5.2034 | -101.2374
5.2071 | -101.2036
5.2109 | -101.1915
5.2146 | -101.1916
5.2183 | -101.2055
5.222 | -101.2292
5.2257 | -101.2568
5.2294 | -101.2796
5.2331 | -101.2855
5.2368 | -101.2679
5.2405 | -101.2368
5.2443 | -101.1918
5.248 | -101.1332
5.2517 | -101.0606
5.2554 | -100.9868
5.2591 | -100.9204
5.2628 | -100.8621
5.2665 | -100.8151
5.2702 | -100.7834
5.274 | -100.76
5.2777 | -100.7498
5.2814 | -100.7544
5.2851 | -100.7704
5.2888 | -100.793
5.2925 | -100.8324
5.2962 | -100.8834
5.2999 | -100.9249
5.3036 | -100.9564
5.3074 | -100.9905
5.3111 | -101.021
5.3148 | -101.0496
5.3185 | -101.069
5.3222 | -101.0919
5.3259 | -101.1205
5.3296 | -101.1502
5.3333 | -101.1828
5.337 | -101.2133
5.3408 | -101.2457
5.3445 | -101.2736
5.3482 | -101.2832
5.3519 | -101.2584
5.3556 | -101.2474
5.3593 | -101.2304
5.363 | -101.2123
5.3667 | -101.1943
5.3704 | -101.1776
5.3742 | -101.1547
5.3779 | -101.1346
5.3816 | -101.1302
5.3853 | -101.1307
5.389 | -101.1295
5.3927 | -101.134
5.3964 | -101.1491
5.4001 | -101.1754
5.4039 | -101.2097
5.4076 | -101.2587
5.4113 | -101.319
5.415 | -101.3727
5.4187 | -101.4268
5.4224 | -101.4749
5.4261 | -101.514
5.4298 | -101.541
5.4335 | -101.5539
5.4373 | -101.5484
5.441 | -101.5288
5.4447 | -101.4987
5.4484 | -101.4623
5.4521 | -101.4326
5.4558 | -101.4074
5.4595 | -101.394
5.4632 | -101.393
5.4669 | -101.3979
5.4707 | -101.4065
5.4744 | -101.4132
5.4781 | -101.4186
5.4818 | -101.4131
5.4855 | -101.3975
5.4892 | -101.366
5.4929 | -101.3233
5.4966 | -101.2721
5.5003 | -101.2264
5.5041 | -101.1818
5.5078 | -101.1416
5.5115 | -101.1098
5.5152 | -101.0901
5.5189 | -101.0787
5.5226 | -101.0761
5.5263 | -101.0796
5.53 | -101.0876
5.5338 | -101.1
5.5375 | -101.1094
5.5412 | -101.1224
5.5449 | -101.1341
5.5486 | -101.1419
5.5523 | -101.156
5.556 | -101.1676
5.5597 | -101.1748
5.5634 | -101.1764
5.5672 | -101.179
5.5709 | -101.1744
5.5746 | -101.1676
5.5783 | -101.1596
5.582 | -101.1464
5.5857 | -101.1316
5.5894 | -101.1137
5.5931 | -101.0994
5.5968 | -101.0836
5.6006 | -101.0687
5.6043 | -101.0563
5.608 | -101.0504
5.6117 | -101.0475
5.6154 | -101.0484
5.6191 | -101.0592
5.6228 | -101.0753
5.6265 | -101.0953
5.6302 | -101.1225
5.634 | -101.1583
5.6377 | -101.2004
5.6414 | -101.2529
5.6451 | -101.3229
5.6488 | -101.4019
5.6525 | -101.4921
5.6562 | -101.5937
5.6599 | -101.6929
5.6637 | -101.7845
5.6674 | -101.8705
5.6711 | -101.9394
5.6748 | -101.9937
5.6785 | -102.0284
5.6822 | -102.0494
5.6859 | -102.0476
5.6896 | -102.033
5.6933 | -102.0063
5.6971 | -101.9733
5.7008 | -101.9391
5.7045 | -101.8987
5.7082 | -101.8634
5.7119 | -101.8328
5.7156 | -101.7938
5.7193 | -101.7547
5.723 | -101.7158
5.7267 | -101.6793
5.7305 | -101.6293
5.7342 | -101.5794
5.7379 | -101.5286
5.7416 | -101.4808
5.7453 | -101.4381
5.749 | -101.4004
5.7527 | -101.3733
5.7564 | -101.3584
5.7601 | -101.3678
5.7639 | -101.3941
5.7676 | -101.4338
5.7713 | -101.4859
5.775 | -101.5432
5.7787 | -101.5977
5.7824 | -101.6453
5.7861 | -101.6702
5.7898 | -101.6735
5.7936 | -101.6534
5.7973 | -101.6145
5.801 | -101.5703
5.8047 | -101.5156
5.8084 | -101.4595
5.8121 | -101.4109
5.8158 | -101.3842
5.8195 | -101.3784
5.8232 | -101.3919
5.827 | -101.4242
5.8307 | -101.4762
5.8344 | -101.545
5.8381 | -101.6253
5.8418 | -101.7072
5.8455 | -101.7849
5.8492 | -101.8684
5.8529 | -101.9464
5.8566 | -102.0158
5.8604 | -102.0659
5.8641 | -102.1105
5.8678 | -102.1393
5.8715 | -102.1484
5.8752 | -102.1488
5.8789 | -102.141
5.8826 | -102.121
5.8863 | -102.0821
5.89 | -102.0551
5.8938 | -102.029
5.8975 | -102.0003
5.9012 | -101.9743
5.9049 | -101.9422
5.9086 | -101.9174
5.9123 | -101.9031
5.916 | -101.8973
5.9197 | -101.8934
5.9235 | -101.8814
5.9272 | -101.869
5.9309 | -101.8616
5.9346 | -101.847
5.9383 | -101.8307
5.942 | -101.8194
5.9457 | -101.8111
5.9494 | -101.814
5.9531 | -101.8179
5.9569 | -101.832
5.9606 | -101.8577
5.9643 | -101.8831
5.968 | -101.9184
5.9717 | -101.9651
5.9754 | -102.0191
5.9791 | -102.0779
5.9828 | -102.1377
5.9865 | -102.1983
5.9903 | -102.2641
5.994 | -102.3192
5.9977 | -102.3709
6.0014 | -102.4064
6.0051 | -102.4315
6.0088 | -102.4479
6.0125 | -102.4565
6.0162 | -102.4606
6.0199 | -102.4543
6.0237 | -102.4438
6.0274 | -102.4388
6.0311 | -102.4316
6.0348 | -102.4269
6.0385 | -102.4314
6.0422 | -102.4405
6.0459 | -102.4573
6.0496 | -102.4734
6.0534 | -102.4877
6.0571 | -102.5083
6.0608 | -102.5228
6.0645 | -102.5376
6.0682 | -102.5554
6.0719 | -102.5711
6.0756 | -102.5905
6.0793 | -102.6083
6.083 | -102.6135
6.0868 | -102.6144
6.0905 | -102.6006
6.0942 | -102.5764
6.0979 | -102.5454
6.1016 | -102.5117
6.1053 | -102.4686
6.109 | -102.4256
6.1127 | -102.3835
6.1164 | -102.3409
6.1202 | -102.3023
6.1239 | -102.27
6.1276 | -102.2483
6.1313 | -102.2346
6.135 | -102.2306
6.1387 | -102.2431
6.1424 | -102.2669
6.1461 | -102.3
6.1498 | -102.3355
6.1536 | -102.3723
6.1573 | -102.3974
6.161 | -102.4141
6.1647 | -102.4242
6.1684 | -102.4126
6.1721 | -102.3892
6.1758 | -102.3625
6.1795 | -102.328
6.1833 | -102.2821
6.187 | -102.2225
6.1907 | -102.1588
6.1944 | -102.0841
6.1981 | -102.0068
6.2018 | -101.9284
6.2055 | -101.852
6.2092 | -101.7842
6.2129 | -101.7244
6.2167 | -101.6767
6.2204 | -101.6397
6.2241 | -101.6165
6.2278 | -101.6106
6.2315 | -101.6184
6.2352 | -101.6388
6.2389 | -101.6627
6.2426 | -101.6907
6.2463 | -101.7237
6.2501 | -101.7612
6.2538 | -101.799
6.2575 | -101.831
6.2612 | -101.8643
6.2649 | -101.8924
6.2686 | -101.9154
6.2723 | -101.9384
6.276 | -101.9679
6.2797 | -101.9965
6.2835 | -102.0236
6.2872 | -102.0619
6.2909 | -102.1072
6.2946 | -102.1597
6.2983 | -102.2208
6.302 | -102.2891
6.3057 | -102.3661
6.3094 | -102.4469
6.3132 | -102.5296
6.3169 | -102.605
6.3206 | -102.6676
6.3243 | -102.7133
6.328 | -102.7446
6.3317 | -102.7653
6.3354 | -102.774
6.3391 | -102.766
6.3428 | -102.7434
6.3466 | -102.7007
6.3503 | -102.6489
6.354 | -102.5884
6.3577 | -102.522
6.3614 | -102.4576
6.3651 | -102.4
6.3688 | -102.3521
6.3725 | -102.3242
6.3762 | -102.3174
6.38 | -102.326
6.3837 | -102.3522
6.3874 | -102.394
6.3911 | -102.4419
6.3948 | -102.493
6.3985 | -102.5482
6.4022 | -102.6026
6.4059 | -102.6522
6.4097 | -102.7003
6.4134 | -102.7394
6.4171 | -102.7725
6.4208 | -102.7984
6.4245 | -102.8127
6.4282 | -102.8296
6.4319 | -102.8331
6.4356 | -102.8378
6.4393 | -102.8343
6.4431 | -102.8327
6.4468 | -102.8219
6.4505 | -102.7991
6.4542 | -102.7685
6.4579 | -102.7173
6.4616 | -102.6532
6.4653 | -102.5871
6.469 | -102.5159
6.4727 | -102.4387
6.4765 | -102.3654
6.4802 | -102.2948
6.4839 | -102.2303
6.4876 | -102.1836
6.4913 | -102.154
6.495 | -102.1327
6.4987 | -102.1306
6.5024 | -102.148
6.5061 | -102.1809
6.5099 | -102.2281
6.5136 | -102.2831
6.5173 | -102.3482
6.521 | -102.4164
6.5247 | -102.4799
6.5284 | -102.547
6.5321 | -102.6066
6.5358 | -102.6474
6.5396 | -102.6786
6.5433 | -102.6929
6.547 | -102.6945
6.5507 | -102.6884
6.5544 | -102.6667
6.5581 | -102.6414
6.5618 | -102.6092
6.5655 | -102.5678
6.5692 | -102.5279
6.573 | -102.4986
6.5767 | -102.4801
6.5804 | -102.4773
6.5841 | -102.4871
6.5878 | -102.5089
6.5915 | -102.5387
6.5952 | -102.584
6.5989 | -102.6361
6.6026 | -102.6883
6.6064 | -102.7433
6.6101 | -102.7929
6.6138 | -102.8338
6.6175 | -102.8681
6.6212 | -102.8942
6.6249 | -102.9108
6.6286 | -102.9151
6.6323 | -102.9206
6.636 | -102.9257
6.6398 | -102.9266
6.6435 | -102.928
6.6472 | -102.9358
6.6509 | -102.9449
6.6546 | -102.9571
6.6583 | -102.9729
6.662 | -102.9885
6.6657 | -103.0036
6.6695 | -103.0183
6.6732 | -103.0289
6.6769 | -103.0328
6.6806 | -103.0265
6.6843 | -103.0166
6.688 | -103.0025
6.6917 | -102.9874
6.6954 | -102.9675
6.6991 | -102.9412
6.7029 | -102.9146
6.7066 | -102.8947
6.7103 | -102.8775
6.714 | -102.8641
6.7177 | -102.849
6.7214 | -102.8417
6.7251 | -102.8366
6.7288 | -102.8293
6.7325 | -102.818
6.7363 | -102.8075
6.74 | -102.7918
6.7437 | -102.7711
6.7474 | -102.7495
6.7511 | -102.7272
6.7548 | -102.7038
6.7585 | -102.6727
6.7622 | -102.6412
6.7659 | -102.6156
6.7697 | -102.5896
6.7734 | -102.5651
6.7771 | -102.5567
6.7808 | -102.5613
6.7845 | -102.5591
6.7882 | -102.5583
6.7919 | -102.5636
6.7956 | -102.5712
6.7993 | -102.5872
6.8031 | -102.6065
6.8068 | -102.6223
6.8105 | -102.6468
6.8142 | -102.6721
6.8179 | -102.6985
6.8216 | -102.7213
6.8253 | -102.7494
6.829 | -102.7801
6.8328 | -102.8123
6.8365 | -102.8445
6.8402 | -102.8775
6.8439 | -102.9044
6.8476 | -102.9272
6.8513 | -102.9461
6.855 | -102.9605
6.8587 | -102.9795
6.8624 | -102.9956
6.8662 | -103.0056
6.8699 | -103.0193
6.8736 | -103.0391
6.8773 | -103.0528
6.881 | -103.0674
6.8847 | -103.0891
6.8884 | -103.1101
6.8921 | -103.1249
6.8958 | -103.1278
6.8996 | -103.1205
6.9033 | -103.1073
6.907 | -103.0893
6.9107 | -103.0675
6.9144 | -103.0426
6.9181 | -103.0187
6.9218 | -103.002
6.9255 | -102.9926
6.9293 | -102.9937
6.933 | -103.0079
6.9367 | -103.0322
6.9404 | -103.0646
6.9441 | -103.1018
6.9478 | -103.1442
6.9515 | -103.1874
6.9552 | -103.2249
6.9589 | -103.199
6.9627 | -103.1838
6.9664 | -103.1533
6.9701 | -103.1114
6.9738 | -103.059
6.9775 | -102.9989
6.9812 | -102.9365
6.9849 | -102.8738
6.9886 | -102.8235
6.9923 | -102.7813
6.9961 | -102.7539
6.9998 | -102.7495
7.0035 | -102.7615
7.0072 | -102.7892
7.0109 | -102.8324
7.0146 | -102.8819
7.0183 | -102.9403
7.022 | -103.0037
7.0257 | -103.0693
7.0295 | -103.1292
7.0332 | -103.1778
7.0369 | -103.2157
7.0406 | -103.2447
7.0443 | -103.2674
7.048 | -103.2787
7.0517 | -103.2703
7.0554 | -103.2469
7.0592 | -103.2169
7.0629 | -103.1898
7.0666 | -103.1643
7.0703 | -103.1479
7.074 | -103.1329
7.0777 | -103.1264
7.0814 | -103.1294
7.0851 | -103.1434
7.0888 | -103.1722
7.0926 | -103.2171
7.0963 | -103.264
7.1 | -103.3131
7.1037 | -103.3673
7.1074 | -103.4242
7.1111 | -103.4788
7.1148 | -103.5243
7.1185 | -103.5615
7.1222 | -103.5972
7.126 | -103.6223
7.1297 | -103.6364
7.1334 | -103.6444
7.1371 | -103.6516
7.1408 | -103.6557
7.1445 | -103.6592
7.1482 | -103.6587
7.1519 | -103.6535
7.1556 | -103.6545
7.1594 | -103.6615
7.1631 | -103.6693
7.1668 | -103.6759
7.1705 | -103.6792
7.1742 | -103.6862
7.1779 | -103.6913
7.1816 | -103.6948
7.1853 | -103.6895
7.1891 | -103.6848
7.1928 | -103.6786
7.1965 | -103.6752
7.2002 | -103.669
7.2039 | -103.6598
7.2076 | -103.6488
7.2113 | -103.6234
7.215 | -103.5856
7.2187 | -103.5487
7.2225 | -103.5154
7.2262 | -103.4745
7.2299 | -103.426
7.2336 | -103.3805
7.2373 | -103.3407
7.241 | -103.3027
7.2447 | -103.2676
7.2484 | -103.2378
7.2521 | -103.2158
7.2559 | -103.206
7.2596 | -103.2048
7.2633 | -103.198
7.267 | -103.1899
7.2707 | -103.1868
7.2744 | -103.1873
7.2781 | -103.1941
7.2818 | -103.2028
7.2855 | -103.2171
7.2893 | -103.2344
7.293 | -103.2574
7.2967 | -103.2829
7.3004 | -103.3137
7.3041 | -103.3381
7.3078 | -103.3667
7.3115 | -103.3969
7.3152 | -103.4245
7.319 | -103.4515
7.3227 | -103.4699
7.3264 | -103.4802
7.3301 | -103.4846
7.3338 | -103.4888
7.3375 | -103.4885
7.3412 | -103.4815
7.3449 | -103.4727
7.3486 | -103.4645
7.3524 | -103.4611
7.3561 | -103.4584
7.3598 | -103.4542
7.3635 | -103.4488
7.3672 | -103.4432
7.3709 | -103.4494
7.3746 | -103.4538
7.3783 | -103.4566
7.382 | -103.4568
7.3858 | -103.4526
7.3895 | -103.4451
7.3932 | -103.4225
7.3969 | -103.3908
7.4006 | -103.3571
7.4043 | -103.3201
7.408 | -103.2768
7.4117 | -103.2371
7.4154 | -103.1956
7.4192 | -103.1554
7.4229 | -103.1222
7.4266 | -103.094
7.4303 | -103.0806
7.434 | -103.074
7.4377 | -103.0792
7.4414 | -103.0934
7.4451 | -103.1178
7.4489 | -103.1512
7.4526 | -103.1874
7.4563 | -103.2256
7.46 | -103.2655
7.4637 | -103.3081
7.4674 | -103.3458
7.4711 | -103.383
7.4748 | -103.4148
7.4785 | -103.438
7.4823 | -103.4626
7.486 | -103.4843
7.4897 | -103.5014
7.4934 | -103.525
7.4971 | -103.5442
7.5008 | -103.5692
7.5045 | -103.604
7.5082 | -103.6419
7.5119 | -103.6706
7.5157 | -103.7006
7.5194 | -103.735
7.5231 | -103.764
7.5268 | -103.7965
7.5305 | -103.8315
7.5342 | -103.8599
7.5379 | -103.8855
7.5416 | -103.889
7.5453 | -103.8787
7.5491 | -103.8527
7.5528 | -103.8188
7.5565 | -103.7744
7.5602 | -103.7206
7.5639 | -103.6594
7.5676 | -103.6007
7.5713 | -103.542
7.575 | -103.4853
7.5788 | -103.4434
7.5825 | -103.4102
7.5862 | -103.3867
7.5899 | -103.3695
7.5936 | -103.361
7.5973 | -103.3641
7.601 | -103.3784
7.6047 | -103.398
7.6084 | -103.4208
7.6122 | -103.4463
7.6159 | -103.4692
7.6196 | -103.4887
7.6233 | -103.5117
7.627 | -103.5353
7.6307 | -103.5477
7.6344 | -103.5592
7.6381 | -103.5722
7.6418 | -103.5765
7.6456 | -103.5737
7.6493 | -103.5656
7.653 | -103.5515
7.6567 | -103.5425
7.6604 | -103.5294
7.6641 | -103.5118
7.6678 | -103.4872
7.6715 | -103.4601
7.6752 | -103.4331
7.679 | -103.4091
7.6827 | -103.3842
7.6864 | -103.3631
7.6901 | -103.3494
7.6938 | -103.3328
7.6975 | -103.3176
7.7012 | -103.3105
7.7049 | -103.3099
7.7087 | -103.3085
7.7124 | -103.3081
7.7161 | -103.3088
7.7198 | -103.3135
7.7235 | -103.3211
7.7272 | -103.3325
7.7309 | -103.345
7.7346 | -103.3597
7.7383 | -103.3705
7.7421 | -103.3792
7.7458 | -103.3964
7.7495 | -103.4111
7.7532 | -103.4234
7.7569 | -103.4338
7.7606 | -103.4414
7.7643 | -103.4455
7.768 | -103.4493
7.7717 | -103.4477
7.7755 | -103.4509
7.7792 | -103.4567
7.7829 | -103.464
7.7866 | -103.4787
7.7903 | -103.4961
7.794 | -103.515
7.7977 | -103.5403
7.8014 | -103.5657
7.8052 | -103.5962
7.8089 | -103.6339
7.8126 | -103.6769
7.8163 | -103.716
7.82 | -103.7546
7.8237 | -103.7914
7.8274 | -103.8224
7.8311 | -103.8513
7.8348 | -103.8775
7.8386 | -103.8911
7.8423 | -103.8902
7.846 | -103.8844
7.8497 | -103.8738
7.8534 | -103.8521
7.8571 | -103.8275
7.8608 | -103.7908
7.8645 | -103.7498
7.8682 | -103.7073
7.872 | -103.6655
7.8757 | -103.6294
7.8794 | -103.6013
7.8831 | -103.5777
7.8868 | -103.5598
7.8905 | -103.5398
7.8942 | -103.5213
7.8979 | -103.5124
7.9016 | -103.5068
7.9054 | -103.5044
7.9091 | -103.5092
7.9128 | -103.5144
7.9165 | -103.5292
7.9202 | -103.5477
7.9239 | -103.5681
7.9276 | -103.5898
7.9313 | -103.6149
7.9351 | -103.6404
7.9388 | -103.6665
7.9425 | -103.6931
7.9462 | -103.714
7.9499 | -103.7286
7.9536 | -103.7487
7.9573 | -103.7742
7.961 | -103.7895
7.9647 | -103.8003
7.9685 | -103.8189
7.9722 | -103.84
7.9759 | -103.8645
7.9796 | -103.891
7.9833 | -103.9176
7.987 | -103.9491
7.9907 | -103.9805
7.9944 | -104.0101
7.9981 | -104.0379
8.0019 | -104.0644
8.0056 | -104.0908
8.0093 | -104.1161
8.013 | -104.1299
8.0167 | -104.136
8.0204 | -104.1381
8.0241 | -104.1317
8.0278 | -104.1207
8.0315 | -104.0755
8.0353 | -104.0497
8.039 | -104.0152
8.0427 | -103.9784
8.0464 | -103.9397
8.0501 | -103.9054
8.0538 | -103.8717
8.0575 | -103.8409
8.0612 | -103.8179
8.0649 | -103.8051
8.0687 | -103.8016
8.0724 | -103.8027
8.0761 | -103.8105
8.0798 | -103.8307
8.0835 | -103.8605
8.0872 | -103.8985
8.0909 | -103.9438
8.0946 | -103.9906
8.0984 | -104.042
8.1021 | -104.0981
8.1058 | -104.1521
8.1095 | -104.2051
8.1132 | -104.256
8.1169 | -104.3103
8.1206 | -104.3623
8.1243 | -104.4062
8.128 | -104.4406
8.1318 | -104.4674
8.1355 | -104.4819
8.1392 | -104.4877
8.1429 | -104.4881
8.1466 | -104.4807
8.1503 | -104.4739
8.154 | -104.4612
8.1577 | -104.4428
8.1614 | -104.4255
8.1652 | -104.4088
8.1689 | -104.3906
8.1726 | -104.3661
8.1763 | -104.3421
8.18 | -104.3204
8.1837 | -104.2948
8.1874 | -104.2711
8.1911 | -104.2499
8.1949 | -104.2251
8.1986 | -104.1978
8.2023 | -104.1732
8.206 | -104.1498
8.2097 | -104.1247
8.2134 | -104.1021
8.2171 | -104.085
8.2208 | -104.069
8.2245 | -104.055
8.2283 | -104.0398
8.232 | -104.026
8.2357 | -104.0134
8.2394 | -104.0022
8.2431 | -103.9919
8.2468 | -103.9861
8.2505 | -103.9844
8.2542 | -103.9829
8.2579 | -103.989
8.2617 | -103.9999
8.2654 | -104.0098
8.2691 | -104.0198
8.2728 | -104.0297
8.2765 | -104.047
8.2802 | -104.0668
8.2839 | -104.0856
8.2876 | -104.1089
8.2913 | -104.1331
8.2951 | -104.1535
8.2988 | -104.1745
8.3025 | -104.1923
8.3062 | -104.2104
8.3099 | -104.224
8.3136 | -104.2356
8.3173 | -104.2424
8.321 | -104.2438
8.3248 | -104.2397
8.3285 | -104.2232
8.3322 | -104.2038
8.3359 | -104.1821
8.3396 | -104.151
8.3433 | -104.1126
8.347 | -104.0712
8.3507 | -104.0296
8.3544 | -103.9877
8.3582 | -103.9455
8.3619 | -103.9001
8.3656 | -103.8587
8.3693 | -103.8228
8.373 | -103.7938
8.3767 | -103.7699
8.3804 | -103.7532
8.3841 | -103.7459
8.3878 | -103.7471
8.3916 | -103.7607
8.3953 | -103.7841
8.399 | -103.8106
8.4027 | -103.8394
8.4064 | -103.8715
8.4101 | -103.9035
8.4138 | -103.9384
8.4175 | -103.9747
8.4212 | -104.0068
8.425 | -104.0315
8.4287 | -104.0533
8.4324 | -104.0763
8.4361 | -104.0952
8.4398 | -104.1124
8.4435 | -104.1273
8.4472 | -104.1412
8.4509 | -104.1571
8.4547 | -104.1688
8.4584 | -104.1806
8.4621 | -104.1988
8.4658 | -104.2185
8.4695 | -104.2379
8.4732 | -104.2626
8.4769 | -104.2875
8.4806 | -104.3129
8.4843 | -104.3405
8.4881 | -104.3663
8.4918 | -104.3914
8.4955 | -104.4127
8.4992 | -104.4374
8.5029 | -104.4577
8.5066 | -104.4692
8.5103 | -104.4834
8.514 | -104.4968
8.5177 | -104.5023
8.5215 | -104.5063
8.5252 | -104.5122
8.5289 | -104.5228
8.5326 | -104.5279
8.5363 | -104.5303
8.54 | -104.5374
8.5437 | -104.5402
8.5474 | -104.5446
8.5511 | -104.5492
8.5549 | -104.5552
8.5586 | -104.5621
8.5623 | -104.5681
8.566 | -104.5556
8.5697 | -104.5518
8.5734 | -104.5464
8.5771 | -104.5352
8.5808 | -104.5195
8.5846 | -104.5017
8.5883 | -104.4776
8.592 | -104.4516
8.5957 | -104.4248
8.5994 | -104.3939
8.6031 | -104.3661
8.6068 | -104.3423
8.6105 | -104.3192
8.6142 | -104.3011
8.618 | -104.2849
8.6217 | -104.2741
8.6254 | -104.2713
8.6291 | -104.2738
8.6328 | -104.2884
8.6365 | -104.308
8.6402 | -104.3309
8.6439 | -104.3645
8.6476 | -104.4031
8.6514 | -104.4437
8.6551 | -104.4852
8.6588 | -104.5309
8.6625 | -104.5764
8.6662 | -104.619
8.6699 | -104.6641
8.6736 | -104.7057
8.6773 | -104.739
8.681 | -104.7669
8.6848 | -104.7901
8.6885 | -104.8118
8.6922 | -104.8326
8.6959 | -104.8501
8.6996 | -104.862
8.7033 | -104.8697
8.707 | -104.8797
8.7107 | -104.8874
8.7145 | -104.8907
8.7182 | -104.8949
8.7219 | -104.8943
8.7256 | -104.8902
8.7293 | -104.8842
8.733 | -104.8723
8.7367 | -104.8584
8.7404 | -104.8484
8.7441 | -104.8366
8.7479 | -104.8247
8.7516 | -104.8137
8.7553 | -104.7958
8.759 | -104.7793
8.7627 | -104.7576
8.7664 | -104.7334
8.7701 | -104.7112
8.7738 | -104.6884
8.7775 | -104.6638
8.7813 | -104.6392
8.785 | -104.6181
8.7887 | -104.6001
8.7924 | -104.585
8.7961 | -104.5744
8.7998 | -104.5639
8.8035 | -104.5559
8.8072 | -104.5496
8.8109 | -104.5376
8.8147 | -104.5227
8.8184 | -104.5105
8.8221 | -104.4946
8.8258 | -104.4773
8.8295 | -104.4573
8.8332 | -104.4344
8.8369 | -104.4117
8.8406 | -104.3883
8.8444 | -104.3588
8.8481 | -104.3291
8.8518 | -104.3003
8.8555 | -104.2762
8.8592 | -104.2568
8.8629 | -104.2442
8.8666 | -104.2338
8.8703 | -104.227
8.874 | -104.2301
8.8778 | -104.2414
8.8815 | -104.2579
8.8852 | -104.2821
8.8889 | -104.3127
8.8926 | -104.3526
8.8963 | -104.3938
8.9 | -104.4352
8.9037 | -104.4751
8.9074 | -104.5145
8.9112 | -104.5522
8.9149 | -104.5812
8.9186 | -104.609
8.9223 | -104.6355
8.926 | -104.6567
8.9297 | -104.679
8.9334 | -104.6927
8.9371 | -104.7028
8.9408 | -104.7117
8.9446 | -104.7217
8.9483 | -104.7258
8.952 | -104.7223
8.9557 | -104.7183
8.9594 | -104.7154
8.9631 | -104.7135
8.9668 | -104.71
8.9705 | -104.706
8.9743 | -104.7034
8.978 | -104.705
8.9817 | -104.7092
8.9854 | -104.7148
8.9891 | -104.7241
8.9928 | -104.7323
8.9965 | -104.7429
9.0002 | -104.7542
9.0039 | -104.7654
9.0077 | -104.7745
9.0114 | -104.7888
9.0151 | -104.8088
9.0188 | -104.8251
9.0225 | -104.8446
9.0262 | -104.866
9.0299 | -104.8862
9.0336 | -104.9126
9.0373 | -104.9395
9.0411 | -104.9623
9.0448 | -104.9852
9.0485 | -105.0063
9.0522 | -105.0219
9.0559 | -105.0364
9.0596 | -105.0516
9.0633 | -105.0658
9.067 | -105.0721
9.0707 | -105.0719
9.0745 | -105.0634
9.0782 | -105.0452
9.0819 | -105.0274
9.0856 | -105.0129
9.0893 | -105.0003
9.093 | -104.9878
9.0967 | -104.9761
9.1004 | -104.9726
9.1042 | -104.9691
9.1079 | -104.9698
9.1116 | -104.9685
9.1153 | -104.9684
9.119 | -104.9687
9.1227 | -104.9711
9.1264 | -104.9783
9.1301 | -104.9873
9.1338 | -105.0005
9.1376 | -105.0126
9.1413 | -105.0256
9.145 | -105.0348
9.1487 | -105.0409
9.1524 | -105.0431
9.1561 | -105.0375
9.1598 | -105.0333
9.1635 | -105.0249
9.1672 | -105.0086
9.171 | -104.9882
9.1747 | -104.965
9.1784 | -104.9365
9.1821 | -104.9009
9.1858 | -104.8645
9.1895 | -104.8267
9.1932 | -104.7878
9.1969 | -104.7536
9.2006 | -104.7214
9.2044 | -104.692
9.2081 | -104.6658
9.2118 | -104.6445
9.2155 | -104.6288
9.2192 | -104.6142
9.2229 | -104.6001
9.2266 | -104.5935
9.2303 | -104.5957
9.2341 | -104.5965
9.2378 | -104.5945
9.2415 | -104.5931
9.2452 | -104.5951
9.2489 | -104.5985
9.2526 | -104.6004
9.2563 | -104.6009
9.26 | -104.6006
9.2637 | -104.5982
9.2675 | -104.5924
9.2712 | -104.5843
9.2749 | -104.5773
9.2786 | -104.5715
9.2823 | -104.5655
9.286 | -104.5596
9.2897 | -104.5594
9.2934 | -104.5631
9.2971 | -104.5704
9.3009 | -104.5808
9.3046 | -104.5978
9.3083 | -104.6177
9.312 | -104.6384
9.3157 | -104.6663
9.3194 | -104.6999
9.3231 | -104.7376
9.3268 | -104.7762
9.3306 | -104.8125
9.3343 | -104.8483
9.338 | -104.8872
9.3417 | -104.9274
9.3454 | -104.9599
9.3491 | -104.9884
9.3528 | -105.0081
9.3565 | -105.018
9.3602 | -105.0268
9.364 | -105.0297
9.3677 | -105.0255
9.3714 | -105.0233
9.3751 | -105.0155
9.3788 | -105.0021
9.3825 | -104.9845
9.3862 | -104.9685
9.3899 | -104.9582
9.3936 | -104.9519
9.3974 | -104.9431
9.4011 | -104.9444
9.4048 | -104.9528
9.4085 | -104.9579
9.4122 | -104.9652
9.4159 | -104.9779
9.4196 | -104.9936
9.4233 | -105.0118
9.427 | -105.0292
9.4308 | -105.0447
9.4345 | -105.0589
9.4382 | -105.075
9.4419 | -105.0833
9.4456 | -105.0957
9.4493 | -105.1101
9.453 | -105.119
9.4567 | -105.1272
9.4604 | -105.1306
9.4642 | -105.1291
9.4679 | -105.1287
9.4716 | -105.1262
9.4753 | -105.1207
9.479 | -105.1113
9.4827 | -105.1062
9.4864 | -105.0959
9.4901 | -105.0846
9.4939 | -105.0742
9.4976 | -105.0627
9.5013 | -105.0526
9.505 | -105.0416
9.5087 | -105.0334
9.5124 | -105.0334
9.5161 | -105.0362
9.5198 | -105.0395
9.5235 | -105.0447
9.5273 | -105.0486
9.531 | -105.0571
9.5347 | -105.07
9.5384 | -105.0792
9.5421 | -105.089
9.5458 | -105.103
9.5495 | -105.1149
9.5532 | -105.1239
9.5569 | -105.1315
9.5607 | -105.1416
9.5644 | -105.1529
9.5681 | -105.1647
9.5718 | -105.1769
9.5755 | -105.1886
9.5792 | -105.2027
9.5829 | -105.2122
9.5866 | -105.221
9.5904 | -105.2293
9.5941 | -105.2303
9.5978 | -105.2274
9.6015 | -105.229
9.6052 | -105.2232
9.6089 | -105.2151
9.6126 | -105.1965
9.6163 | -105.1718
9.62 | -105.141
9.6238 | -105.1095
9.6275 | -105.0767
9.6312 | -105.0412
9.6349 | -104.9997
9.6386 | -104.9628
9.6423 | -104.9268
9.646 | -104.8918
9.6497 | -104.8606
9.6534 | -104.8332
9.6572 | -104.8085
9.6609 | -104.7923
9.6646 | -104.7831
9.6683 | -104.7831
9.672 | -104.7879
9.6757 | -104.7985
9.6794 | -104.8119
9.6831 | -104.8294
9.6868 | -104.8539
9.6906 | -104.8836
9.6943 | -104.9126
9.698 | -104.9418
9.7017 | -104.9654
9.7054 | -104.9865
9.7091 | -105.0051
9.7128 | -105.0189
9.7165 | -105.0262
9.7203 | -105.0284
9.724 | -105.0249
9.7277 | -105.0147
9.7314 | -104.9963
9.7351 | -104.9765
9.7388 | -104.9555
9.7425 | -104.927
9.7462 | -104.8988
9.7499 | -104.8639
9.7537 | -104.8364
9.7574 | -104.8101
9.7611 | -104.7857
9.7648 | -104.7667
9.7685 | -104.751
9.7722 | -104.7407
9.7759 | -104.732
9.7796 | -104.7292
9.7833 | -104.7378
9.7871 | -104.7508
9.7908 | -104.7656
9.7945 | -104.784
9.7982 | -104.808
9.8019 | -104.8375
9.8056 | -104.869
9.8093 | -104.8982
9.813 | -104.9279
9.8167 | -104.9602
9.8205 | -104.9936
9.8242 | -105.0276
9.8279 | -105.0572
9.8316 | -105.0816
9.8353 | -105.0998
9.839 | -105.1124
9.8427 | -105.1205
9.8464 | -105.125
9.8502 | -105.1285
9.8539 | -105.1265
9.8576 | -105.1201
9.8613 | -105.1124
9.865 | -105.0994
9.8687 | -105.0855
9.8724 | -105.0691
9.8761 | -105.0505
9.8798 | -105.0316
9.8836 | -105.0173
9.8873 | -105.0032
9.891 | -104.9916
9.8947 | -104.9881
9.8984 | -104.9838
9.9021 | -104.9851
9.9058 | -104.9916
9.9095 | -105.0012
9.9132 | -105.0142
9.917 | -105.0313
9.9207 | -105.055
9.9244 | -105.082
9.9281 | -105.1082
9.9318 | -105.1364
9.9355 | -105.1689
9.9392 | -105.1997
9.9429 | -105.2303
9.9466 | -105.2589
9.9504 | -105.2845
9.9541 | -105.3083
9.9578 | -105.3255
9.9615 | -105.3354
9.9652 | -105.343
9.9689 | -105.3433
9.9726 | -105.3379
9.9763 | -105.327
9.9801 | -105.3127
9.9838 | -105.2977
9.9875 | -105.2776
9.9912 | -105.2539
9.9949 | -105.2218
9.9986 | -105.1905
10.0023 | -105.1651
10.006 | -105.1418
10.0097 | -105.1187
10.0135 | -105.093
10.0172 | -105.0751
10.0209 | -105.0654
10.0246 | -105.0585
10.0283 | -105.0508
10.032 | -105.045
10.0357 | -105.0435
10.0394 | -105.0434
10.0431 | -105.0505
10.0469 | -105.0609
10.0506 | -105.072
10.0543 | -105.0835
10.058 | -105.0974
10.0617 | -105.1157
10.0654 | -105.1354
10.0691 | -105.1511
10.0728 | -105.1618
10.0765 | -105.1774
10.0803 | -105.1883
10.084 | -105.1948
10.0877 | -105.1987
10.0914 | -105.2027
10.0951 | -105.2048
10.0988 | -105.2039
10.1025 | -105.2015
10.1062 | -105.1935
10.11 | -105.1791
10.1137 | -105.163
10.1174 | -105.148
10.1211 | -105.1292
10.1248 | -105.1062
10.1285 | -105.0814
10.1322 | -105.0556
10.1359 | -105.0288
10.1396 | -105.0038
10.1434 | -104.9778
10.1471 | -104.9498
10.1508 | -104.9252
10.1545 | -104.9082
10.1582 | -104.8969
10.1619 | -104.8852
10.1656 | -104.8766
10.1693 | -104.8713
10.173 | -104.8794
10.1768 | -104.8774
10.1805 | -104.8783
10.1842 | -104.8821
10.1879 | -104.8873
10.1916 | -104.8892
10.1953 | -104.8903
10.199 | -104.8892
10.2027 | -104.8904
10.2064 | -104.8937
10.2102 | -104.8912
10.2139 | -104.8898
10.2176 | -104.8886
10.2213 | -104.8883
10.225 | -104.8867
10.2287 | -104.8893
10.2324 | -104.8911
10.2361 | -104.8938
10.2399 | -104.8963
10.2436 | -104.9016
10.2473 | -104.9101
10.251 | -104.9167
10.2547 | -104.923
10.2584 | -104.9336
10.2621 | -104.9467
10.2658 | -104.9618
10.2695 | -104.9827
10.2733 | -105.0078
10.277 | -105.0334
10.2807 | -105.0606
10.2844 | -105.0889
10.2881 | -105.1201
10.2918 | -105.1491
10.2955 | -105.1782
10.2992 | -105.2054
10.3029 | -105.2323
10.3067 | -105.2567
10.3104 | -105.2804
10.3141 | -105.3028
10.3178 | -105.3241
10.3215 | -105.3393
10.3252 | -105.3552
10.3289 | -105.3714
10.3326 | -105.3854
10.3363 | -105.3919
10.3401 | -105.3949
10.3438 | -105.398
10.3475 | -105.3963
10.3512 | -105.3928
10.3549 | -105.3852
10.3586 | -105.3704
10.3623 | -105.351
10.366 | -105.3287
10.3698 | -105.3116
10.3735 | -105.2925
10.3772 | -105.2687
10.3809 | -105.2459
10.3846 | -105.2239
10.3883 | -105.2016
10.392 | -105.1845
10.3957 | -105.1674
10.3994 | -105.1548
10.4032 | -105.1478
10.4069 | -105.145
10.4106 | -105.1418
10.4143 | -105.1423
10.418 | -105.1433
10.4217 | -105.1447
10.4254 | -105.1497
10.4291 | -105.1582
10.4328 | -105.1685
10.4366 | -105.1814
10.4403 | -105.1914
10.444 | -105.2028
10.4477 | -105.216
10.4514 | -105.2315
10.4551 | -105.2427
10.4588 | -105.2557
10.4625 | -105.271
10.4662 | -105.2843
10.47 | -105.2994
10.4737 | -105.3178
10.4774 | -105.3352
10.4811 | -105.3491
10.4848 | -105.3625
10.4885 | -105.3787
10.4922 | -105.3971
10.4959 | -105.4126
10.4997 | -105.4296
10.5034 | -105.4438
10.5071 | -105.4517
10.5108 | -105.4585
10.5145 | -105.4666
10.5182 | -105.4744
10.5219 | -105.4747
10.5256 | -105.4698
10.5293 | -105.4659
10.5331 | -105.4589
10.5368 | -105.4502
10.5405 | -105.4413
10.5442 | -105.4302
10.5479 | -105.4182
10.5516 | -105.4015
10.5553 | -105.3815
10.559 | -105.3616
10.5627 | -105.3408
10.5665 | -105.3124
10.5702 | -105.2862
10.5739 | -105.263
10.5776 | -105.2421
10.5813 | -105.2203
10.585 | -105.2
10.5887 | -105.1785
10.5924 | -105.1613
10.5961 | -105.1465
10.5999 | -105.1284
10.6036 | -105.115
10.6073 | -105.11
10.611 | -105.1079
10.6147 | -105.1052
10.6184 | -105.1069
10.6221 | -105.109
10.6258 | -105.1104
10.6296 | -105.1133
10.6333 | -105.1167
10.637 | -105.1251
10.6407 | -105.1318
10.6444 | -105.1329
10.6481 | -105.1369
10.6518 | -105.1408
10.6555 | -105.1432
10.6592 | -105.1404
10.663 | -105.1364
10.6667 | -105.1289
10.6704 | -105.1138
10.6741 | -105.0949
10.6778 | -105.0816
10.6815 | -105.0664
10.6852 | -105.0495
10.6889 | -105.0332
10.6926 | -105.0188
10.6964 | -105.0065
10.7001 | -104.9947
10.7038 | -104.9916
10.7075 | -105.0093
10.7112 | -105.0191
10.7149 | -105.0331
10.7186 | -105.049
10.7223 | -105.074
10.7261 | -105.1015
10.7298 | -105.1283
10.7335 | -105.157
10.7372 | -105.187
10.7409 | -105.2174
10.7446 | -105.2468
10.7483 | -105.2732
10.752 | -105.2962
10.7557 | -105.3209
10.7595 | -105.338
10.7632 | -105.3462
10.7669 | -105.3512
10.7706 | -105.3476
10.7743 | -105.338
10.778 | -105.3308
10.7817 | -105.3243
10.7854 | -105.314
10.7891 | -105.3021
10.7929 | -105.2858
10.7966 | -105.2731
10.8003 | -105.2638
10.804 | -105.2555
10.8077 | -105.2468
10.8114 | -105.2412
10.8151 | -105.2439
10.8188 | -105.2474
10.8225 | -105.2567
10.8263 | -105.2688
10.83 | -105.2868
10.8337 | -105.3077
10.8374 | -105.332
10.8411 | -105.3582
10.8448 | -105.3864
10.8485 | -105.4167
10.8522 | -105.4447
10.8559 | -105.4757
10.8597 | -105.5054
10.8634 | -105.5322
10.8671 | -105.5494
10.8708 | -105.563
10.8745 | -105.5729
10.8782 | -105.5815
10.8819 | -105.5815
10.8856 | -105.5767
10.8894 | -105.5731
10.8931 | -105.562
10.8968 | -105.5486
10.9005 | -105.5309
10.9042 | -105.5113
10.9079 | -105.4931
10.9116 | -105.4736
10.9153 | -105.451
10.919 | -105.4277
10.9228 | -105.4078
10.9265 | -105.3901
10.9302 | -105.3798
10.9339 | -105.3713
10.9376 | -105.3568
10.9413 | -105.3446
10.945 | -105.3334
10.9487 | -105.3248
10.9524 | -105.3183
10.9562 | -105.3091
10.9599 | -105.3053
10.9636 | -105.3041
10.9673 | -105.3078
10.971 | -105.3114
10.9747 | -105.3133
10.9784 | -105.3155
10.9821 | -105.3139
10.9858 | -105.3141
10.9896 | -105.3172
10.9933 | -105.3162
10.997 | -105.3139
11.0007 | -105.3176
11.0044 | -105.3239
11.0081 | -105.3279
11.0118 | -105.3347
11.0155 | -105.3426
11.0193 | -105.3468
11.023 | -105.3528
11.0267 | -105.3659
11.0304 | -105.3745
11.0341 | -105.3851
11.0378 | -105.3969
11.0415 | -105.4151
11.0452 | -105.4351
11.0489 | -105.4507
11.0527 | -105.4683
11.0564 | -105.4879
11.0601 | -105.5067
11.0638 | -105.5249
11.0675 | -105.5439
11.0712 | -105.5595
11.0749 | -105.5779
11.0786 | -105.5997
11.0823 | -105.6134
11.0861 | -105.6228
11.0898 | -105.6299
11.0935 | -105.6329
11.0972 | -105.6342
11.1009 | -105.6292
11.1046 | -105.6227
11.1083 | -105.6079
11.112 | -105.5884
11.1158 | -105.57
11.1195 | -105.5477
11.1232 | -105.52
11.1269 | -105.4913
11.1306 | -105.463
11.1343 | -105.4369
11.138 | -105.4099
11.1417 | -105.3854
11.1454 | -105.3645
11.1492 | -105.3445
11.1529 | -105.3272
11.1566 | -105.3144
11.1603 | -105.3082
11.164 | -105.2994
11.1677 | -105.2925
11.1714 | -105.2924
11.1751 | -105.2945
11.1788 | -105.3014
11.1826 | -105.3134
11.1863 | -105.3214
11.19 | -105.3317
11.1937 | -105.3432
11.1974 | -105.3547
11.2011 | -105.3662
11.2048 | -105.3719
11.2085 | -105.3784
11.2122 | -105.3862
11.216 | -105.3897
11.2197 | -105.3911
11.2234 | -105.3871
11.2271 | -105.3756
11.2308 | -105.3613
11.2345 | -105.3362
11.2382 | -105.3089
11.2419 | -105.2715
11.2457 | -105.2461
11.2494 | -105.2168
11.2531 | -105.1884
11.2568 | -105.1614
11.2605 | -105.1356
11.2642 | -105.1091
11.2679 | -105.0839
11.2716 | -105.0635
11.2753 | -105.0441
11.2791 | -105.0256
11.2828 | -105.0081
11.2865 | -104.9938
11.2902 | -104.9827
11.2939 | -104.9747
11.2976 | -104.9671
11.3013 | -104.9682
11.305 | -104.974
11.3087 | -104.9809
11.3125 | -104.9865
11.3162 | -104.997
11.3199 | -105.0101
11.3236 | -105.024
11.3273 | -105.0447
11.331 | -105.068
11.3347 | -105.0898
11.3384 | -105.1137
11.3421 | -105.1359
11.3459 | -105.1539
11.3496 | -105.1757
11.3533 | -105.1943
11.357 | -105.2075
11.3607 | -105.2235
11.3644 | -105.2378
11.3681 | -105.2536
11.3718 | -105.2664
11.3756 | -105.2741
11.3793 | -105.281
11.383 | -105.2889
11.3867 | -105.2984
11.3904 | -105.3072
11.3941 | -105.3173
11.3978 | -105.33
11.4015 | -105.3392
11.4052 | -105.3475
11.409 | -105.3535
11.4127 | -105.3577
11.4164 | -105.3662
11.4201 | -105.3709
11.4238 | -105.3782
11.4275 | -105.386
11.4312 | -105.3955
11.4349 | -105.4058
11.4386 | -105.4153
11.4424 | -105.4244
11.4461 | -105.4334
11.4498 | -105.444
11.4535 | -105.4568
11.4572 | -105.4681
11.4609 | -105.478
11.4646 | -105.4869
11.4683 | -105.4946
11.472 | -105.5009
11.4758 | -105.5047
11.4795 | -105.5147
11.4832 | -105.5252
11.4869 | -105.5352
11.4906 | -105.543
11.4943 | -105.5497
11.498 | -105.5594
11.5017 | -105.5694
11.5055 | -105.5778
11.5092 | -105.5913
11.5129 | -105.6066
11.5166 | -105.6174
11.5203 | -105.6281
11.524 | -105.6381
11.5277 | -105.6421
11.5314 | -105.6427
11.5351 | -105.6399
11.5389 | -105.6371
11.5426 | -105.632
11.5463 | -105.6242
11.55 | -105.6098
11.5537 | -105.5892
11.5574 | -105.5685
11.5611 | -105.5472
11.5648 | -105.5239
11.5685 | -105.4976
11.5723 | -105.4693
11.576 | -105.4416
11.5797 | -105.4138
11.5834 | -105.3811
11.5871 | -105.3514
11.5908 | -105.3216
11.5945 | -105.2922
11.5982 | -105.2666
11.6019 | -105.2433
11.6057 | -105.2223
11.6094 | -105.2031
11.6131 | -105.1872
11.6168 | -105.1742
11.6205 | -105.164
11.6242 | -105.1588
11.6279 | -105.1569
11.6316 | -105.1577
11.6354 | -105.1571
11.6391 | -105.162
11.6428 | -105.1738
11.6465 | -105.1857
11.6502 | -105.1978
11.6539 | -105.2118
11.6576 | -105.2301
11.6613 | -105.25
11.665 | -105.2702
11.6688 | -105.2925
11.6725 | -105.3135
11.6762 | -105.3322
11.6799 | -105.3557
11.6836 | -105.3817
11.6873 | -105.4063
11.691 | -105.4318
11.6947 | -105.4562
11.6984 | -105.4793
11.7022 | -105.5027
11.7059 | -105.5222
11.7096 | -105.5384
11.7133 | -105.5543
11.717 | -105.5713
11.7207 | -105.5863
11.7244 | -105.5975
11.7281 | -105.6036
11.7318 | -105.6099
11.7356 | -105.6191
11.7393 | -105.627
11.743 | -105.6312
11.7467 | -105.6307
11.7504 | -105.626
11.7541 | -105.6218
11.7578 | -105.616
11.7615 | -105.6093
11.7653 | -105.6017
11.769 | -105.5936
11.7727 | -105.5883
11.7764 | -105.5784
11.7801 | -105.5724
11.7838 | -105.5643
11.7875 | -105.5572
11.7912 | -105.5524
11.7949 | -105.5467
11.7987 | -105.5428
11.8024 | -105.5409
11.8061 | -105.5412
11.8098 | -105.5432
11.8135 | -105.5453
11.8172 | -105.5508
11.8209 | -105.5602
11.8246 | -105.5728
11.8283 | -105.5836
11.8321 | -105.5953
11.8358 | -105.6069
11.8395 | -105.6206
11.8432 | -105.6391
11.8469 | -105.6573
11.8506 | -105.6718
11.8543 | -105.6875
11.858 | -105.7057
11.8617 | -105.7181
11.8655 | -105.7268
11.8692 | -105.7352
11.8729 | -105.7446
11.8766 | -105.7503
11.8803 | -105.7539
11.884 | -105.7564
11.8877 | -105.7555
11.8914 | -105.7548
11.8952 | -105.7516
11.8989 | -105.7461
11.9026 | -105.7381
11.9063 | -105.725
11.91 | -105.7088
11.9137 | -105.6925
11.9174 | -105.6769
11.9211 | -105.6633
11.9248 | -105.6449
11.9286 | -105.6261
11.9323 | -105.6105
11.936 | -105.5921
11.9397 | -105.5784
11.9434 | -105.5647
11.9471 | -105.5555
11.9508 | -105.5459
11.9545 | -105.5383
11.9582 | -105.5342
11.962 | -105.5306
11.9657 | -105.5277
11.9694 | -105.5258
11.9731 | -105.5262
11.9768 | -105.5265
11.9805 | -105.5291
11.9842 | -105.5293
11.9879 | -105.5313
11.9916 | -105.5363
11.9954 | -105.5426
11.9991 | -105.549
12.0028 | -105.5605
12.0065 | -105.5769
12.0102 | -105.5918
12.0139 | -105.6033
12.0176 | -105.6151
12.0213 | -105.6299
12.0251 | -105.6447
12.0288 | -105.6607
12.0325 | -105.6775
12.0362 | -105.698
12.0399 | -105.718
12.0436 | -105.734
12.0473 | -105.7413
12.051 | -105.7498
12.0547 | -105.7606
12.0585 | -105.7661
12.0622 | -105.7709
12.0659 | -105.7752
12.0696 | -105.7785
12.0733 | -105.7821
12.077 | -105.7833
12.0807 | -105.7776
12.0844 | -105.7639
12.0881 | -105.7568
12.0919 | -105.7461
12.0956 | -105.7289
12.0993 | -105.7093
12.103 | -105.6931
12.1067 | -105.6731
12.1104 | -105.6545
12.1141 | -105.6296
12.1178 | -105.6061
12.1215 | -105.5837
12.1253 | -105.561
12.129 | -105.5361
12.1327 | -105.5116
12.1364 | -105.4913
12.1401 | -105.4724
12.1438 | -105.459
12.1475 | -105.4478
12.1512 | -105.4343
12.155 | -105.4245
12.1587 | -105.4204
12.1624 | -105.4217
12.1661 | -105.4252
12.1698 | -105.4301
12.1735 | -105.4373
12.1772 | -105.4426
12.1809 | -105.4495
12.1846 | -105.4622
12.1884 | -105.4772
12.1921 | -105.4911
12.1958 | -105.5044
12.1995 | -105.5177
12.2032 | -105.5299
12.2069 | -105.5382
12.2106 | -105.5402
12.2143 | -105.542
12.218 | -105.5442
12.2218 | -105.5461
12.2255 | -105.5415
12.2292 | -105.5312
12.2329 | -105.5192
12.2366 | -105.508
12.2403 | -105.4944
12.244 | -105.4778
12.2477 | -105.4609
12.2515 | -105.4469
12.2552 | -105.4399
12.2589 | -105.4397
12.2626 | -105.4271
12.2663 | -105.4172
12.27 | -105.4125
12.2737 | -105.4083
12.2774 | -105.3994
12.2811 | -105.3953
12.2849 | -105.3944
12.2886 | -105.3969
12.2923 | -105.3987
12.296 | -105.4034
12.2997 | -105.4104
12.3034 | -105.4211
12.3071 | -105.4308
12.3108 | -105.439
12.3145 | -105.4547
12.3183 | -105.4673
12.322 | -105.4786
12.3257 | -105.4908
12.3294 | -105.5092
12.3331 | -105.5211
12.3368 | -105.532
12.3405 | -105.5441
12.3442 | -105.5567
12.3479 | -105.5716
12.3517 | -105.5871
12.3554 | -105.6043
12.3591 | -105.6217
12.3628 | -105.6379
12.3665 | -105.6516
12.3702 | -105.6615
12.3739 | -105.6694
12.3776 | -105.677
12.3813 | -105.6835
12.3851 | -105.69
12.3888 | -105.6969
12.3925 | -105.7024
12.3962 | -105.7036
12.3999 | -105.7065
12.4036 | -105.7083
12.4073 | -105.7088
12.411 | -105.7096
12.4148 | -105.7058
12.4185 | -105.7021
12.4222 | -105.7017
12.4259 | -105.7031
12.4296 | -105.7002
12.4333 | -105.6966
12.437 | -105.6927
12.4407 | -105.6924
12.4444 | -105.6932
12.4482 | -105.6924
12.4519 | -105.6918
12.4556 | -105.6902
12.4593 | -105.6845
12.463 | -105.6759
12.4667 | -105.6686
12.4704 | -105.6645
12.4741 | -105.6572
12.4778 | -105.6508
12.4816 | -105.6471
12.4853 | -105.6448
12.489 | -105.6409
12.4927 | -105.6356
12.4964 | -105.6334
12.5001 | -105.6325
12.5038 | -105.634
12.5075 | -105.6387
12.5113 | -105.641
12.515 | -105.6482
12.5187 | -105.654
12.5224 | -105.6585
12.5261 | -105.6638
12.5298 | -105.6707
12.5335 | -105.6802
12.5372 | -105.6873
12.5409 | -105.6935
12.5447 | -105.6996
12.5484 | -105.7098
12.5521 | -105.7236
12.5558 | -105.734
12.5595 | -105.7449
12.5632 | -105.755
12.5669 | -105.7655
12.5706 | -105.7765
12.5743 | -105.7881
12.5781 | -105.8
12.5818 | -105.8089
12.5855 | -105.8187
12.5892 | -105.8274
12.5929 | -105.8374
12.5966 | -105.8506
12.6003 | -105.8616
12.604 | -105.8726
12.6077 | -105.8798
12.6115 | -105.8908
12.6152 | -105.9033
12.6189 | -105.9156
12.6226 | -105.927
12.6263 | -105.9389
12.63 | -105.9514
12.6337 | -105.967
12.6374 | -105.9831
12.6412 | -106.0021
12.6449 | -106.0251
12.6486 | -106.0453
12.6523 | -106.064
12.656 | -106.0839
12.6597 | -106.1057
12.6634 | -106.1244
12.6671 | -106.1415
12.6708 | -106.1617
12.6746 | -106.1821
12.6783 | -106.1957
12.682 | -106.212
12.6857 | -106.2223
12.6894 | -106.2271
12.6931 | -106.2356
12.6968 | -106.2412
12.7005 | -106.247
12.7042 | -106.2469
12.708 | -106.2434
12.7117 | -106.2388
12.7154 | -106.2306
12.7191 | -106.2214
12.7228 | -106.2082
12.7265 | -106.1915
12.7302 | -106.1745
12.7339 | -106.1597
12.7376 | -106.1459
12.7414 | -106.1298
12.7451 | -106.1125
12.7488 | -106.0962
12.7525 | -106.0799
12.7562 | -106.0624
12.7599 | -106.0469
12.7636 | -106.0315
12.7673 | -106.0173
12.7711 | -106.0043
12.7748 | -105.9894
12.7785 | -105.9747
12.7822 | -105.9609
12.7859 | -105.9462
12.7896 | -105.9327
12.7933 | -105.9182
12.797 | -105.9013
12.8007 | -105.8844
12.8045 | -105.8715
12.8082 | -105.8529
12.8119 | -105.8352
12.8156 | -105.8189
12.8193 | -105.8073
12.823 | -105.7945
12.8267 | -105.7787
12.8304 | -105.7647
12.8341 | -105.7515
12.8379 | -105.7387
12.8416 | -105.722
12.8453 | -105.7069
12.849 | -105.6956
12.8527 | -105.6835
12.8564 | -105.6724
12.8601 | -105.6627
12.8638 | -105.6534
12.8675 | -105.6442
12.8713 | -105.6372
12.875 | -105.6307
12.8787 | -105.6245
12.8824 | -105.619
12.8861 | -105.6155
12.8898 | -105.6146
12.8935 | -105.6134
12.8972 | -105.6123
12.901 | -105.615
12.9047 | -105.6219
12.9084 | -105.6315
12.9121 | -105.6405
12.9158 | -105.6494
12.9195 | -105.6557
12.9232 | -105.661
12.9269 | -105.6699
12.9306 | -105.6789
12.9344 | -105.6866
12.9381 | -105.6934
12.9418 | -105.698
12.9455 | -105.6994
12.9492 | -105.7025
12.9529 | -105.7053
12.9566 | -105.7049
12.9603 | -105.7037
12.964 | -105.7009
12.9678 | -105.6966
12.9715 | -105.6903
12.9752 | -105.6842
12.9789 | -105.6757
12.9826 | -105.6685
12.9863 | -105.662
12.99 | -105.6534
12.9937 | -105.6477
12.9974 | -105.6402
13.0012 | -105.6294
13.0049 | -105.6188
13.0086 | -105.6069
13.0123 | -105.5982
13.016 | -105.589
13.0197 | -105.5818
13.0234 | -105.5764
13.0271 | -105.5765
13.0309 | -105.5782
13.0346 | -105.5834
13.0383 | -105.5887
13.042 | -105.5934
13.0457 | -105.5967
13.0494 | -105.6014
13.0531 | -105.6115
13.0568 | -105.621
13.0605 | -105.6322
13.0643 | -105.6436
13.068 | -105.6513
13.0717 | -105.6584
13.0754 | -105.6604
13.0791 | -105.6639
13.0828 | -105.67
13.0865 | -105.673
13.0902 | -105.672
13.0939 | -105.6706
13.0977 | -105.6662
13.1014 | -105.6637
13.1051 | -105.6594
13.1088 | -105.6574
13.1125 | -105.6562
13.1162 | -105.6522
13.1199 | -105.6486
13.1236 | -105.6448
13.1273 | -105.6399
13.1311 | -105.637
13.1348 | -105.6354
13.1385 | -105.6303
13.1422 | -105.6277
13.1459 | -105.6293
13.1496 | -105.6314
13.1533 | -105.6349
13.157 | -105.6396
13.1608 | -105.6452
13.1645 | -105.6479
13.1682 | -105.6515
13.1719 | -105.6568
13.1756 | -105.6614
13.1793 | -105.667
13.183 | -105.6737
13.1867 | -105.6782
13.1904 | -105.6794
13.1942 | -105.6787
13.1979 | -105.6788
13.2016 | -105.675
13.2053 | -105.67
13.209 | -105.6645
13.2127 | -105.6583
13.2164 | -105.6472
13.2201 | -105.6351
13.2238 | -105.621
13.2276 | -105.6104
13.2313 | -105.6027
13.235 | -105.5944
13.2387 | -105.5835
13.2424 | -105.5688
13.2461 | -105.556
13.2498 | -105.5456
13.2535 | -105.5365
13.2572 | -105.529
13.261 | -105.523
13.2647 | -105.5194
13.2684 | -105.5131
13.2721 | -105.5049
13.2758 | -105.4985
13.2795 | -105.4907
13.2832 | -105.4811
13.2869 | -105.4755
13.2907 | -105.4722
13.2944 | -105.468
13.2981 | -105.4612
13.3018 | -105.4574
13.3055 | -105.454
13.3092 | -105.4504
13.3129 | -105.4471
13.3166 | -105.4451
13.3203 | -105.4424
13.3241 | -105.44
13.3278 | -105.44
13.3315 | -105.4345
13.3352 | -105.4262
13.3389 | -105.4226
13.3426 | -105.4196
13.3463 | -105.4173
13.35 | -105.4164
13.3537 | -105.4167
13.3575 | -105.4207
13.3612 | -105.4235
13.3649 | -105.4253
13.3686 | -105.4295
13.3723 | -105.4373
13.376 | -105.4487
13.3797 | -105.4633
13.3834 | -105.4794
13.3871 | -105.4948
13.3909 | -105.5094
13.3946 | -105.5254
13.3983 | -105.5343
13.402 | -105.5438
13.4057 | -105.5535
13.4094 | -105.5639
13.4131 | -105.5701
13.4168 | -105.5744
13.4206 | -105.5787
13.4243 | -105.579
13.428 | -105.5776
13.4317 | -105.576
13.4354 | -105.572
13.4391 | -105.5667
13.4428 | -105.5613
13.4465 | -105.5539
13.4502 | -105.5457
13.454 | -105.5361
13.4577 | -105.5286
13.4614 | -105.5241
13.4651 | -105.523
13.4688 | -105.5214
13.4725 | -105.5226
13.4762 | -105.5273
13.4799 | -105.5326
13.4836 | -105.5363
13.4874 | -105.5432
13.4911 | -105.5548
13.4948 | -105.5654
13.4985 | -105.575
13.5022 | -105.5879
13.5059 | -105.6029
13.5096 | -105.621
13.5133 | -105.6352
13.517 | -105.6483
13.5208 | -105.6639
13.5245 | -105.6834
13.5282 | -105.7022
13.5319 | -105.7155
13.5356 | -105.7248
13.5393 | -105.7329
13.543 | -105.7374
13.5467 | -105.7441
13.5505 | -105.7449
13.5542 | -105.7448
13.5579 | -105.7435
13.5616 | -105.7376
13.5653 | -105.7299
13.569 | -105.7202
13.5727 | -105.7117
13.5764 | -105.702
13.5801 | -105.6882
13.5839 | -105.6726
13.5876 | -105.6593
13.5913 | -105.6469
13.595 | -105.638
13.5987 | -105.6233
13.6024 | -105.6048
13.6061 | -105.586
13.6098 | -105.5709
13.6135 | -105.5558
13.6173 | -105.5409
13.621 | -105.5271
13.6247 | -105.5175
13.6284 | -105.5065
13.6321 | -105.4968
13.6358 | -105.4885
13.6395 | -105.4825
13.6432 | -105.4813
13.647 | -105.4794
13.6507 | -105.4793
13.6544 | -105.4814
13.6581 | -105.4838
13.6618 | -105.4861
13.6655 | -105.4904
13.6692 | -105.4959
13.6729 | -105.4977
13.6766 | -105.4983
13.6804 | -105.5035
13.6841 | -105.5094
13.6878 | -105.5138
13.6915 | -105.5188
13.6952 | -105.5228
13.6989 | -105.5264
13.7026 | -105.5253
13.7063 | -105.5219
13.71 | -105.5203
13.7138 | -105.5175
13.7175 | -105.5113
13.7212 | -105.503
13.7249 | -105.4952
13.7286 | -105.4911
13.7323 | -105.4892
13.736 | -105.4856
13.7397 | -105.4847
13.7434 | -105.4821
13.7472 | -105.4784
13.7509 | -105.4746
13.7546 | -105.4699
13.7583 | -105.4667
13.762 | -105.4628
13.7657 | -105.459
13.7694 | -105.4538
13.7731 | -105.4503
13.7768 | -105.4493
13.7806 | -105.4475
13.7843 | -105.4448
13.788 | -105.4423
13.7917 | -105.4408
13.7954 | -105.4413
13.7991 | -105.4407
13.8028 | -105.4412
13.8065 | -105.443
13.8103 | -105.445
13.814 | -105.4469
13.8177 | -105.4496
13.8214 | -105.4499
13.8251 | -105.4505
13.8288 | -105.4519
13.8325 | -105.4518
13.8362 | -105.452
13.8399 | -105.4502
13.8437 | -105.4541
13.8474 | -105.4589
13.8511 | -105.4613
13.8548 | -105.4648
13.8585 | -105.472
13.8622 | -105.4796
13.8659 | -105.489
13.8696 | -105.4996
13.8733 | -105.5126
13.8771 | -105.5262
13.8808 | -105.5402
13.8845 | -105.5571
13.8882 | -105.5762
13.8919 | -105.5965
13.8956 | -105.6124
13.8993 | -105.6345
13.903 | -105.6608
13.9068 | -105.6885
13.9105 | -105.7179
13.9142 | -105.7465
13.9179 | -105.7735
13.9216 | -105.8029
13.9253 | -105.829
13.929 | -105.8492
13.9327 | -105.8706
13.9364 | -105.8933
13.9402 | -105.9104
13.9439 | -105.9231
13.9476 | -105.9368
13.9513 | -105.9476
13.955 | -105.9541
13.9587 | -105.9321
13.9624 | -105.8711
13.9661 | -105.8299
13.9698 | -105.7955
13.9736 | -105.7504
13.9773 | -105.6973
13.981 | -105.6428
13.9847 | -105.5801
13.9884 | -105.5104
13.9921 | -105.4355
13.9958 | -105.3517
13.9995 | -105.2659
14.0032 | -105.1753
14.007 | -105.0794
14.0107 | -104.984
14.0144 | -104.8892
14.0181 | -104.7944
14.0218 | -104.7028
14.0255 | -104.6162
14.0292 | -104.5301
14.0329 | -104.4475
14.0366 | -104.3712
14.0404 | -104.3039
14.0441 | -104.242
14.0478 | -104.1898
14.0515 | -104.1482
14.0552 | -104.1143
14.0589 | -104.0885
14.0626 | -104.0732
14.0663 | -104.0664
14.0701 | -104.0705
14.0738 | -104.0834
14.0775 | -104.1026
14.0812 | -104.1299
14.0849 | -104.1648
14.0886 | -104.2069
14.0923 | -104.2585
14.096 | -104.3163
14.0997 | -104.3768
14.1035 | -104.4376
14.1072 | -104.5007
14.1109 | -104.565
14.1146 | -104.6291
14.1183 | -104.6934
14.122 | -104.756
14.1257 | -104.8179
14.1294 | -104.88
14.1331 | -104.9412
14.1369 | -104.997
14.1406 | -105.049
14.1443 | -105.0999
14.148 | -105.1482
14.1517 | -105.1912
14.1554 | -105.2279
14.1591 | -105.2634
14.1628 | -105.3109
14.1666 | -105.3728
14.1703 | -105.4006
14.174 | -105.4109
14.1777 | -105.4232
14.1814 | -105.4349
14.1851 | -105.4448
14.1888 | -105.4543
14.1925 | -105.4624
14.1962 | -105.4709
14.2 | -105.4803
14.2037 | -105.4885
14.2074 | -105.4957
14.2111 | -105.5011
14.2148 | -105.5061
14.2185 | -105.5116
14.2222 | -105.5165
14.2259 | -105.5204
14.2296 | -105.5236
14.2334 | -105.5287
14.2371 | -105.5345
14.2408 | -105.5393
14.2445 | -105.5417
14.2482 | -105.5431
14.2519 | -105.5407
14.2556 | -105.5389
14.2593 | -105.5395
14.263 | -105.5416
14.2668 | -105.5452
14.2705 | -105.5449
14.2742 | -105.5446
14.2779 | -105.5479
14.2816 | -105.5553
14.2853 | -105.5641
14.289 | -105.5725
14.2927 | -105.5817
14.2965 | -105.5884
14.3002 | -105.5969
14.3039 | -105.6042
14.3076 | -105.6121
14.3113 | -105.6236
14.315 | -105.6378
14.3187 | -105.6533
14.3224 | -105.6667
14.3261 | -105.6759
14.3299 | -105.6827
14.3336 | -105.6899
14.3373 | -105.6966
14.341 | -105.704
14.3447 | -105.7111
14.3484 | -105.7141
14.3521 | -105.7179
14.3558 | -105.7238
14.3595 | -105.7311
14.3633 | -105.7335
14.367 | -105.7337
14.3707 | -105.7326
14.3744 | -105.729
14.3781 | -105.7247
14.3818 | -105.7205
14.3855 | -105.7133
14.3892 | -105.7021
14.3929 | -105.6892
14.3967 | -105.6771
14.4004 | -105.6643
14.4041 | -105.652
14.4078 | -105.6389
14.4115 | -105.6246
14.4152 | -105.611
14.4189 | -105.5955
14.4226 | -105.5788
14.4264 | -105.5624
14.4301 | -105.5437
14.4338 | -105.5262
14.4375 | -105.5086
14.4412 | -105.4905
14.4449 | -105.4712
14.4486 | -105.4526
14.4523 | -105.4357
14.456 | -105.4197
14.4598 | -105.4043
14.4635 | -105.387
14.4672 | -105.3686
14.4709 | -105.3497
14.4746 | -105.3305
14.4783 | -105.3165
14.482 | -105.3003
14.4857 | -105.2813
14.4894 | -105.2627
14.4932 | -105.2446
14.4969 | -105.2305
14.5006 | -105.2176
14.5043 | -105.2061
14.508 | -105.1999
14.5117 | -105.1957
14.5154 | -105.1943
14.5191 | -105.1931
14.5228 | -105.1914
14.5266 | -105.1915
14.5303 | -105.1971
14.534 | -105.2064
14.5377 | -105.2172
14.5414 | -105.2274
14.5451 | -105.2378
14.5488 | -105.2488
14.5525 | -105.263
14.5563 | -105.2786
14.56 | -105.2939
14.5637 | -105.3098
14.5674 | -105.3273
14.5711 | -105.3486
14.5748 | -105.3684
14.5785 | -105.3872
14.5822 | -105.408
14.5859 | -105.4282
14.5897 | -105.4472
14.5934 | -105.4634
14.5971 | -105.4786
14.6008 | -105.4955
14.6045 | -105.5123
14.6082 | -105.5274
14.6119 | -105.5388
14.6156 | -105.5468
14.6193 | -105.5514
14.6231 | -105.5565
14.6268 | -105.5595
14.6305 | -105.5629
14.6342 | -105.5653
14.6379 | -105.5661
14.6416 | -105.564
14.6453 | -105.5603
14.649 | -105.5522
14.6527 | -105.5407
14.6565 | -105.5318
14.6602 | -105.5205
14.6639 | -105.5051
14.6676 | -105.4895
14.6713 | -105.4755
14.675 | -105.4615
14.6787 | -105.4471
14.6824 | -105.4287
14.6862 | -105.4095
14.6899 | -105.3899
14.6936 | -105.3726
14.6973 | -105.3543
14.701 | -105.3366
14.7047 | -105.3164
14.7084 | -105.2954
14.7121 | -105.277
14.7158 | -105.2592
14.7196 | -105.2383
14.7233 | -105.2182
14.727 | -105.1983
14.7307 | -105.1809
14.7344 | -105.1669
14.7381 | -105.155
14.7418 | -105.1389
14.7455 | -105.1257
14.7492 | -105.1146
14.753 | -105.1075
14.7567 | -105.1006
14.7604 | -105.0951
14.7641 | -105.0888
14.7678 | -105.0855
14.7715 | -105.0841
14.7752 | -105.081
14.7789 | -105.0826
14.7826 | -105.0844
14.7864 | -105.0874
14.7901 | -105.0902
14.7938 | -105.0914
14.7975 | -105.0972
14.8012 | -105.1026
14.8049 | -105.1113
14.8086 | -105.1211
14.8123 | -105.1284
14.8161 | -105.1346
14.8198 | -105.1392
14.8235 | -105.1444
14.8272 | -105.1508
14.8309 | -105.1585
14.8346 | -105.1647
14.8383 | -105.1716
14.842 | -105.1753
14.8457 | -105.1788
14.8495 | -105.1851
14.8532 | -105.1902
14.8569 | -105.1965
14.8606 | -105.2042
14.8643 | -105.2086
14.868 | -105.2085
14.8717 | -105.2125
14.8754 | -105.2151
14.8791 | -105.2182
14.8829 | -105.2199
14.8866 | -105.2198
14.8903 | -105.2197
14.894 | -105.2198
14.8977 | -105.221
14.9014 | -105.2216
14.9051 | -105.2207
14.9088 | -105.2198
14.9125 | -105.2188
14.9163 | -105.2204
14.92 | -105.2187
14.9237 | -105.2158
14.9274 | -105.216
14.9311 | -105.2149
14.9348 | -105.2144
14.9385 | -105.2157
14.9422 | -105.2166
14.946 | -105.2173
14.9497 | -105.2172
14.9534 | -105.2214
14.9571 | -105.2227
14.9608 | -105.2246
14.9645 | -105.2287
14.9682 | -105.2344
14.9719 | -105.2381
14.9756 | -105.2415
14.9794 | -105.2474
14.9831 | -105.2535
14.9868 | -105.258
14.9905 | -105.2263
14.9942 | -105.2201
14.9979 | -105.215
15.0016 | -105.2122
15.0053 | -105.2074
15.009 | -105.2006
15.0128 | -105.1902
15.0165 | -105.1792
15.0202 | -105.1685
15.0239 | -105.1574
15.0276 | -105.1436
15.0313 | -105.1273
15.035 | -105.1112
15.0387 | -105.0962
15.0425 | -105.0778
15.0462 | -105.0597
15.0499 | -105.0394
15.0536 | -105.0196
15.0573 | -105.0021
15.061 | -104.9833
15.0647 | -104.9647
15.0684 | -104.9456
15.0721 | -104.9284
15.0759 | -104.9141
15.0796 | -104.9021
15.0833 | -104.8906
15.087 | -104.8794
15.0907 | -104.8715
15.0944 | -104.8665
15.0981 | -104.8626
15.1018 | -104.8596
15.1055 | -104.8587
15.1093 | -104.8622
15.113 | -104.8657
15.1167 | -104.8695
15.1204 | -104.8752
15.1241 | -104.8824
15.1278 | -104.8895
15.1315 | -104.8992
15.1352 | -104.9116
15.1389 | -104.9221
15.1427 | -104.9326
15.1464 | -104.9458
15.1501 | -104.9598
15.1538 | -104.9723
15.1575 | -104.9829
15.1612 | -104.9942
15.1649 | -105.0065
15.1686 | -105.0203
15.1723 | -105.0356
15.1761 | -105.0511
15.1798 | -105.0647
15.1835 | -105.0785
15.1872 | -105.0943
15.1909 | -105.11
15.1946 | -105.123
15.1983 | -105.1354
15.202 | -105.1532
15.2058 | -105.1859
15.2095 | -105.2113
15.2132 | -105.2262
15.2169 | -105.2395
15.2206 | -105.2535
15.2243 | -105.269
15.228 | -105.2803
15.2317 | -105.2941
15.2354 | -105.3082
15.2392 | -105.3225
15.2429 | -105.3347
15.2466 | -105.3468
15.2503 | -105.3592
15.254 | -105.3691
15.2577 | -105.3709
15.2614 | -105.3738
15.2651 | -105.3735
15.2688 | -105.3706
15.2726 | -105.3637
15.2763 | -105.3492
15.28 | -105.336
15.2837 | -105.3216
15.2874 | -105.307
15.2911 | -105.286
15.2948 | -105.2621
15.2985 | -105.2398
15.3023 | -105.217
15.306 | -105.192
15.3097 | -105.1649
15.3134 | -105.1344
15.3171 | -105.1032
15.3208 | -105.0755
15.3245 | -105.0454
15.3282 | -105.0147
15.3319 | -104.984
15.3357 | -104.9532
15.3394 | -104.9249
15.3431 | -104.8985
15.3468 | -104.8763
15.3505 | -104.8512
15.3542 | -104.8268
15.3579 | -104.8097
15.3616 | -104.7939
15.3653 | -104.7793
15.3691 | -104.7679
15.3728 | -104.7593
15.3765 | -104.7539
15.3802 | -104.7511
15.3839 | -104.751
15.3876 | -104.7515
15.3913 | -104.7534
15.395 | -104.7574
15.3987 | -104.7631
15.4025 | -104.7683
15.4062 | -104.772
15.4099 | -104.7794
15.4136 | -104.7882
15.4173 | -104.7977
15.421 | -104.8069
15.4247 | -104.8157
15.4284 | -104.8225
15.4322 | -104.8253
15.4359 | -104.828
15.4396 | -104.8321
15.4433 | -104.8369
15.447 | -104.8426
15.4507 | -104.8433
15.4544 | -104.841
15.4581 | -104.835
15.4618 | -104.8271
15.4656 | -104.8215
15.4693 | -104.8155
15.473 | -104.8103
15.4767 | -104.8038
15.4804 | -104.7966
15.4841 | -104.7902
15.4878 | -104.7826
15.4915 | -104.7727
15.4952 | -104.7658
15.499 | -104.7569
15.5027 | -104.7452
15.5064 | -104.7333
15.5101 | -104.7232
15.5138 | -104.7126
15.5175 | -104.7021
15.5212 | -104.6924
15.5249 | -104.6901
15.5286 | -104.6853
15.5324 | -104.6818
15.5361 | -104.6792
15.5398 | -104.6788
15.5435 | -104.6795
15.5472 | -104.684
15.5509 | -104.6865
15.5546 | -104.69
15.5583 | -104.692
15.5621 | -104.6951
15.5658 | -104.7
15.5695 | -104.7065
15.5732 | -104.7167
15.5769 | -104.7237
15.5806 | -104.7246
15.5843 | -104.7257
15.588 | -104.7307
15.5917 | -104.7334
15.5955 | -104.7342
15.5992 | -104.7319
15.6029 | -104.7301
15.6066 | -104.7278
15.6103 | -104.7249
15.614 | -104.7206
15.6177 | -104.7151
15.6214 | -104.7093
15.6251 | -104.7051
15.6289 | -104.6996
15.6326 | -104.6916
15.6363 | -104.6824
15.64 | -104.6712
15.6437 | -104.6584
15.6474 | -104.645
15.6511 | -104.635
15.6548 | -104.6278
15.6585 | -104.6193
15.6623 | -104.6086
15.666 | -104.5979
15.6697 | -104.5896
15.6734 | -104.5844
15.6771 | -104.5814
15.6808 | -104.5809
15.6845 | -104.581
15.6882 | -104.5807
15.692 | -104.578
15.6957 | -104.5744
15.6994 | -104.5772
15.7031 | -104.5788
15.7068 | -104.5816
15.7105 | -104.5878
15.7142 | -104.5961
15.7179 | -104.6057
15.7216 | -104.6147
15.7254 | -104.6256
15.7291 | -104.6388
15.7328 | -104.6509
15.7365 | -104.6651
15.7402 | -104.6813
15.7439 | -104.6974
15.7476 | -104.7139
15.7513 | -104.7286
15.755 | -104.7406
15.7588 | -104.7542
15.7625 | -104.771
15.7662 | -104.7872
15.7699 | -104.8016
15.7736 | -104.8145
15.7773 | -104.8298
15.781 | -104.8456
15.7847 | -104.8625
15.7884 | -104.8744
15.7922 | -104.8849
15.7959 | -104.8927
15.7996 | -104.899
15.8033 | -104.9092
15.807 | -104.9213
15.8107 | -104.9292
15.8144 | -104.9316
15.8181 | -104.9354
15.8219 | -104.9396
15.8256 | -104.9466
15.8293 | -104.9493
15.833 | -104.9484
15.8367 | -104.945
15.8404 | -104.946
15.8441 | -104.9469
15.8478 | -104.9448
15.8515 | -104.942
15.8553 | -104.94
15.859 | -104.9355
15.8627 | -104.9312
15.8664 | -104.9297
15.8701 | -104.9278
15.8738 | -104.9263
15.8775 | -104.9215
15.8812 | -104.9165
15.8849 | -104.9131
15.8887 | -104.908
15.8924 | -104.9013
15.8961 | -104.8954
15.8998 | -104.8871
15.9035 | -104.88
15.9072 | -104.8719
15.9109 | -104.8648
15.9146 | -104.8588
15.9183 | -104.8529
15.9221 | -104.8481
15.9258 | -104.8415
15.9295 | -104.8376
15.9332 | -104.8327
15.9369 | -104.8258
15.9406 | -104.8185
15.9443 | -104.8121
15.948 | -104.807
15.9518 | -104.8031
15.9555 | -104.8009
15.9592 | -104.8021
15.9629 | -104.8006
15.9666 | -104.7949
15.9703 | -104.7901
15.974 | -104.7901
15.9777 | -104.792
15.9814 | -104.7927
15.9852 | -104.7945
15.9889 | -104.7953
15.9926 | -104.7942
15.9963 | -104.7965
16 | -104.7996
16.0037 | -104.803
16.0074 | -104.8052
16.0111 | -104.8082
16.0148 | -104.8117
16.0186 | -104.8135
16.0223 | -104.8143
16.026 | -104.8164
16.0297 | -104.816
16.0334 | -104.8155
16.0371 | -104.8189
16.0408 | -104.8213
16.0445 | -104.8209
16.0482 | -104.8208
16.052 | -104.8193
16.0557 | -104.8196
16.0594 | -104.82
16.0631 | -104.8136
16.0668 | -104.813
16.0705 | -104.8077
16.0742 | -104.7999
16.0779 | -104.7955
16.0817 | -104.7911
16.0854 | -104.7847
16.0891 | -104.778
16.0928 | -104.7694
16.0965 | -104.7558
16.1002 | -104.7447
16.1039 | -104.7366
16.1076 | -104.7281
16.1113 | -104.7192
16.1151 | -104.7093
16.1188 | -104.6996
16.1225 | -104.6928
16.1262 | -104.6867
16.1299 | -104.6808
16.1336 | -104.6751
16.1373 | -104.673
16.141 | -104.6727
16.1447 | -104.671
16.1485 | -104.6726
16.1522 | -104.6768
16.1559 | -104.6822
16.1596 | -104.6864
16.1633 | -104.6926
16.167 | -104.7014
16.1707 | -104.7112
16.1744 | -104.722
16.1781 | -104.7338
16.1819 | -104.744
16.1856 | -104.7542
16.1893 | -104.7658
16.193 | -104.7803
16.1967 | -104.7965
16.2004 | -104.812
16.2041 | -104.8248
16.2078 | -104.8382
16.2116 | -104.8536
16.2153 | -104.8676
16.219 | -104.8809
16.2227 | -104.8966
16.2264 | -104.9106
16.2301 | -104.9257
16.2338 | -104.9388
16.2375 | -104.9528
16.2412 | -104.9662
16.245 | -104.9796
16.2487 | -104.9926
16.2524 | -105.0049
16.2561 | -105.015
16.2598 | -105.0256
16.2635 | -105.0337
16.2672 | -105.0389
16.2709 | -105.0461
16.2746 | -105.0529
16.2784 | -105.0584
16.2821 | -105.0641
16.2858 | -105.0655
16.2895 | -105.0658
16.2932 | -105.0689
16.2969 | -105.0705
16.3006 | -105.0728
16.3043 | -105.0762
16.308 | -105.0822
16.3118 | -105.0869
16.3155 | -105.0915
16.3192 | -105.099
16.3229 | -105.1043
16.3266 | -105.1093
16.3303 | -105.1201
16.334 | -105.1272
16.3377 | -105.1309
16.3415 | -105.1385
16.3452 | -105.1476
16.3489 | -105.1584
16.3526 | -105.1696
16.3563 | -105.1785
16.36 | -105.188
16.3637 | -105.1972
16.3674 | -105.2071
16.3711 | -105.2144
16.3749 | -105.222
16.3786 | -105.2312
16.3823 | -105.2397
16.386 | -105.2469
16.3897 | -105.2536
16.3934 | -105.2597
16.3971 | -105.2653
16.4008 | -105.2716
16.4045 | -105.2803
16.4083 | -105.2869
16.412 | -105.2917
16.4157 | -105.295
16.4194 | -105.2974
16.4231 | -105.3001
16.4268 | -105.2994
16.4305 | -105.2975
16.4342 | -105.2979
16.4379 | -105.2963
16.4417 | -105.2951
16.4454 | -105.2902
16.4491 | -105.2852
16.4528 | -105.2782
16.4565 | -105.2702
16.4602 | -105.2627
16.4639 | -105.2547
16.4676 | -105.2492
16.4714 | -105.2415
16.4751 | -105.2328
16.4788 | -105.2221
16.4825 | -105.2132
16.4862 | -105.2042
16.4899 | -105.1932
16.4936 | -105.1834
16.4973 | -105.1769
16.501 | -105.1697
16.5048 | -105.1603
16.5085 | -105.1512
16.5122 | -105.1415
16.5159 | -105.128
16.5196 | -105.1205
16.5233 | -105.1135
16.527 | -105.1046
16.5307 | -105.0957
16.5344 | -105.0848
16.5382 | -105.0736
16.5419 | -105.063
16.5456 | -105.052
16.5493 | -105.0412
16.553 | -105.0297
16.5567 | -105.0196
16.5604 | -105.015
16.5641 | -105.0089
16.5678 | -105.0044
16.5716 | -105.0009
16.5753 | -104.9961
16.579 | -104.993
16.5827 | -104.989
16.5864 | -104.9853
16.5901 | -104.982
16.5938 | -104.9786
16.5975 | -104.9839
16.6013 | -104.9849
16.605 | -104.9879
16.6087 | -104.9925
16.6124 | -104.9988
16.6161 | -105.0019
16.6198 | -105.0054
16.6235 | -105.0126
16.6272 | -105.019
16.6309 | -105.0268
16.6347 | -105.0356
16.6384 | -105.044
16.6421 | -105.048
16.6458 | -105.0483
16.6495 | -105.0523
16.6532 | -105.0581
16.6569 | -105.0612
16.6606 | -105.0623
16.6643 | -105.0635
16.6681 | -105.0648
16.6718 | -105.0633
16.6755 | -105.0635
16.6792 | -105.0616
16.6829 | -105.061
16.6866 | -105.0611
16.6903 | -105.0626
16.694 | -105.0639
16.6977 | -105.0622
16.7015 | -105.0599
16.7052 | -105.0594
16.7089 | -105.0575
16.7126 | -105.0543
16.7163 | -105.0553
16.72 | -105.0555
16.7237 | -105.0566
16.7274 | -105.0615
16.7312 | -105.0643
16.7349 | -105.065
16.7386 | -105.0671
16.7423 | -105.0728
16.746 | -105.0788
16.7497 | -105.0875
16.7534 | -105.0997
16.7571 | -105.1081
16.7608 | -105.1186
16.7646 | -105.1271
16.7683 | -105.1369
16.772 | -105.1515
16.7757 | -105.1631
16.7794 | -105.1727
16.7831 | -105.1839
16.7868 | -105.196
16.7905 | -105.2075
16.7942 | -105.2181
16.798 | -105.2248
16.8017 | -105.2288
16.8054 | -105.2335
16.8091 | -105.2371
16.8128 | -105.2402
16.8165 | -105.2408
16.8202 | -105.2411
16.8239 | -105.2417
16.8276 | -105.2418
16.8314 | -105.238
16.8351 | -105.2324
16.8388 | -105.2247
16.8425 | -105.2158
16.8462 | -105.2033
16.8499 | -105.1916
16.8536 | -105.1772
16.8573 | -105.1671
16.8611 | -105.1576
16.8648 | -105.1429
16.8685 | -105.1282
16.8722 | -105.1151
16.8759 | -105.1015
16.8796 | -105.0885
16.8833 | -105.0771
16.887 | -105.0693
16.8907 | -105.0575
16.8945 | -105.0474
16.8982 | -105.0379
16.9019 | -105.0275
16.9056 | -105.0159
16.9093 | -105.0041
16.913 | -104.9954
16.9167 | -104.9869
16.9204 | -104.9783
16.9241 | -104.9696
16.9279 | -104.958
16.9316 | -104.945
16.9353 | -104.9339
16.939 | -104.9243
16.9427 | -104.9157
16.9464 | -104.9028
16.9501 | -104.8951
16.9538 | -104.8873
16.9576 | -104.8755
16.9613 | -104.8631
16.965 | -104.8523
16.9687 | -104.8386
16.9724 | -104.8252
16.9761 | -104.8102
16.9798 | -104.7924
16.9835 | -104.7768
16.9872 | -104.7634
16.991 | -104.7479
16.9947 | -104.732
16.9984 | -104.7184
17.0021 | -104.7071
17.0058 | -104.6984
17.0095 | -104.6882
17.0132 | -104.6771
17.0169 | -104.6714
17.0206 | -104.6662
17.0244 | -104.6584
17.0281 | -104.6514
17.0318 | -104.6491
17.0355 | -104.6494
17.0392 | -104.6504
17.0429 | -104.6558
17.0466 | -104.6608
17.0503 | -104.6675
17.054 | -104.6756
17.0578 | -104.6843
17.0615 | -104.6928
17.0652 | -104.7034
17.0689 | -104.7162
17.0726 | -104.7293
17.0763 | -104.7395
17.08 | -104.7513
17.0837 | -104.7641
17.0875 | -104.7768
17.0912 | -104.7888
17.0949 | -104.8016
17.0986 | -104.8126
17.1023 | -104.8216
17.106 | -104.8319
17.1097 | -104.8393
17.1134 | -104.8445
17.1171 | -104.8497
17.1209 | -104.8539
17.1246 | -104.8585
17.1283 | -104.8601
17.132 | -104.8549
17.1357 | -104.8568
17.1394 | -104.8556
17.1431 | -104.854
17.1468 | -104.8533
17.1505 | -104.854
17.1543 | -104.8528
17.158 | -104.8475
17.1617 | -104.8413
17.1654 | -104.835
17.1691 | -104.83
17.1728 | -104.8261
17.1765 | -104.8239
17.1802 | -104.819
17.1839 | -104.8143
17.1877 | -104.809
17.1914 | -104.8058
17.1951 | -104.8001
17.1988 | -104.796
17.2025 | -104.7932
17.2062 | -104.7898
17.2099 | -104.785
17.2136 | -104.7822
17.2174 | -104.7796
17.2211 | -104.7779
17.2248 | -104.7784
17.2285 | -104.777
17.2322 | -104.7749
17.2359 | -104.7741
17.2396 | -104.7733
17.2433 | -104.7705
17.247 | -104.7657
17.2508 | -104.763
17.2545 | -104.7603
17.2582 | -104.7578
17.2619 | -104.7558
17.2656 | -104.7556
17.2693 | -104.7586
17.273 | -104.7612
17.2767 | -104.7617
17.2804 | -104.7615
17.2842 | -104.7602
17.2879 | -104.7565
17.2916 | -104.7553
17.2953 | -104.7542
17.299 | -104.7521
17.3027 | -104.7518
17.3064 | -104.7529
17.3101 | -104.7524
17.3138 | -104.7498
17.3176 | -104.7476
17.3213 | -104.7446
17.325 | -104.7396
17.3287 | -104.7347
17.3324 | -104.7305
17.3361 | -104.7239
17.3398 | -104.7167
17.3435 | -104.708
17.3473 | -104.6981
17.351 | -104.6877
17.3547 | -104.6793
17.3584 | -104.6699
17.3621 | -104.6573
17.3658 | -104.6438
17.3695 | -104.6285
17.3732 | -104.6104
17.3769 | -104.5901
17.3807 | -104.5681
17.3844 | -104.5475
17.3881 | -104.528
17.3918 | -104.5073
17.3955 | -104.486
17.3992 | -104.4625
17.4029 | -104.4395
17.4066 | -104.4153
17.4103 | -104.39
17.4141 | -104.3656
17.4178 | -104.3414
17.4215 | -104.3176
17.4252 | -104.2924
17.4289 | -104.2689
17.4326 | -104.2447
17.4363 | -104.2221
17.44 | -104.2021
17.4437 | -104.1829
17.4475 | -104.1628
17.4512 | -104.1396
17.4549 | -104.1201
17.4586 | -104.1025
17.4623 | -104.0867
17.466 | -104.0735
17.4697 | -104.0618
17.4734 | -104.0479
17.4772 | -104.0352
17.4809 | -104.0244
17.4846 | -104.0157
17.4883 | -104.0065
17.492 | -103.9993
17.4957 | -103.9951
17.4994 | -103.9889
17.5031 | -103.9852
17.5068 | -103.9813
17.5106 | -103.9773
17.5143 | -103.9737
17.518 | -103.9715
17.5217 | -103.9694
17.5254 | -103.9689
17.5291 | -103.9684
17.5328 | -103.9693
17.5365 | -103.9707
17.5402 | -103.9718
17.544 | -103.974
17.5477 | -103.9757
17.5514 | -103.9765
17.5551 | -103.9789
17.5588 | -103.9814
17.5625 | -103.984
17.5662 | -103.9861
17.5699 | -103.9894
17.5736 | -103.9961
17.5774 | -103.9989
17.5811 | -104.002
17.5848 | -104.0058
17.5885 | -104.0071
17.5922 | -104.01
17.5959 | -104.0107
17.5996 | -104.0104
17.6033 | -104.0102
17.6071 | -104.007
17.6108 | -104.004
17.6145 | -104.0013
17.6182 | -103.9972
17.6219 | -103.9953
17.6256 | -103.9923
17.6293 | -103.9872
17.633 | -103.979
17.6367 | -103.9708
17.6405 | -103.9626
17.6442 | -103.9528
17.6479 | -103.944
17.6516 | -103.9345
17.6553 | -103.9268
17.659 | -103.9185
17.6627 | -103.9102
17.6664 | -103.9072
17.6701 | -103.8993
17.6739 | -103.8926
17.6776 | -103.8882
17.6813 | -103.8817
17.685 | -103.8772
17.6887 | -103.8726
17.6924 | -103.8695
17.6961 | -103.8682
17.6998 | -103.8698
17.7035 | -103.8747
17.7073 | -103.8757
17.711 | -103.8776
17.7147 | -103.8792
17.7184 | -103.8814
17.7221 | -103.8854
17.7258 | -103.8932
17.7295 | -103.8989
17.7332 | -103.9063
17.737 | -103.913
17.7407 | -103.9227
17.7444 | -103.9314
17.7481 | -103.9374
17.7518 | -103.9443
17.7555 | -103.9523
17.7592 | -103.9613
17.7629 | -103.9673
17.7666 | -103.9721
17.7704 | -103.9768
17.7741 | -103.9813
17.7778 | -103.9846
17.7815 | -103.988
17.7852 | -103.9908
17.7889 | -103.994
17.7926 | -103.9969
17.7963 | -104.0001
17.8 | -104.0038
17.8038 | -104.0053
17.8075 | -104.0044
17.8112 | -104.0041
17.8149 | -104.0052
17.8186 | -104.0082
17.8223 | -104.0099
17.826 | -104.01
17.8297 | -104.0119
17.8334 | -104.0123
17.8372 | -104.0141
17.8409 | -104.0179
17.8446 | -104.0197
17.8483 | -104.0227
17.852 | -104.0239
17.8557 | -104.0248
17.8594 | -104.0297
17.8631 | -104.0333
17.8669 | -104.0341
17.8706 | -104.0355
17.8743 | -104.038
17.878 | -104.0426
17.8817 | -104.0458
17.8854 | -104.0483
17.8891 | -104.05
17.8928 | -104.0517
17.8965 | -104.0534
17.9003 | -104.0551
17.904 | -104.055
17.9077 | -104.0517
17.9114 | -104.0499
17.9151 | -104.0467
17.9188 | -104.0412
17.9225 | -104.0368
17.9262 | -104.0317
17.9299 | -104.0253
17.9337 | -104.0199
17.9374 | -104.0137
17.9411 | -104.0055
17.9448 | -103.9964
17.9485 | -103.9846
17.9522 | -103.9729
17.9559 | -103.9596
17.9596 | -103.942
17.9633 | -103.9286
17.9671 | -103.9179
17.9708 | -103.9067
17.9745 | -103.8938
17.9782 | -103.8783
17.9819 | -103.8634
17.9856 | -103.849
17.9893 | -103.8349
17.993 | -103.8211
17.9968 | -103.8067
18.0005 | -103.7918
18.0042 | -103.7788
18.0079 | -103.7657
18.0116 | -103.7518
18.0153 | -103.7369
18.019 | -103.7252
18.0227 | -103.7161
18.0264 | -103.7061
18.0302 | -103.6974
18.0339 | -103.6891
18.0376 | -103.6813
18.0413 | -103.6749
18.045 | -103.6684
18.0487 | -103.6626
18.0524 | -103.6565
18.0561 | -103.6519
18.0598 | -103.6509
18.0636 | -103.6503
18.0673 | -103.6502
18.071 | -103.6501
18.0747 | -103.6502
18.0784 | -103.6513
18.0821 | -103.6521
18.0858 | -103.6529
18.0895 | -103.6521
18.0933 | -103.6503
18.097 | -103.649
18.1007 | -103.6476
18.1044 | -103.6484
18.1081 | -103.6491
18.1118 | -103.6473
18.1155 | -103.6436
18.1192 | -103.642
18.1229 | -103.6415
18.1267 | -103.6403
18.1304 | -103.6352
18.1341 | -103.6294
18.1378 | -103.6218
18.1415 | -103.6125
18.1452 | -103.6013
18.1489 | -103.5893
18.1526 | -103.5763
18.1563 | -103.5636
18.1601 | -103.5518
18.1638 | -103.5394
18.1675 | -103.524
18.1712 | -103.5095
18.1749 | -103.4942
18.1786 | -103.4775
18.1823 | -103.4601
18.186 | -103.4435
18.1897 | -103.4248
18.1935 | -103.4074
18.1972 | -103.3906
18.2009 | -103.374
18.2046 | -103.3585
18.2083 | -103.3427
18.212 | -103.3274
18.2157 | -103.313
18.2194 | -103.3005
18.2232 | -103.2856
18.2269 | -103.2688
18.2306 | -103.2511
18.2343 | -103.2353
18.238 | -103.2233
18.2417 | -103.2125
18.2454 | -103.2017
18.2491 | -103.1914
18.2528 | -103.1817
18.2566 | -103.1738
18.2603 | -103.1681
18.264 | -103.16
18.2677 | -103.1522
18.2714 | -103.1453
18.2751 | -103.1401
18.2788 | -103.1365
18.2825 | -103.1325
18.2862 | -103.1299
18.29 | -103.1276
18.2937 | -103.1245
18.2974 | -103.1188
18.3011 | -103.1145
18.3048 | -103.1107
18.3085 | -103.1064
18.3122 | -103.1024
18.3159 | -103.0988
18.3196 | -103.0951
18.3234 | -103.0917
18.3271 | -103.0878
18.3308 | -103.0834
18.3345 | -103.0792
18.3382 | -103.0775
18.3419 | -103.0758
18.3456 | -103.0756
18.3493 | -103.0774
18.353 | -103.0785
18.3568 | -103.078
18.3605 | -103.0787
18.3642 | -103.0752
18.3679 | -103.0728
18.3716 | -103.074
18.3753 | -103.0742
18.379 | -103.0719
18.3827 | -103.07
18.3865 | -103.0689
18.3902 | -103.0709
18.3939 | -103.0728
18.3976 | -103.0747
18.4013 | -103.079
18.405 | -103.0839
18.4087 | -103.0902
18.4124 | -103.0943
18.4161 | -103.0973
18.4199 | -103.1011
18.4236 | -103.1049
18.4273 | -103.1103
18.431 | -103.1149
18.4347 | -103.1199
18.4384 | -103.1258
18.4421 | -103.132
18.4458 | -103.1373
18.4495 | -103.1414
18.4533 | -103.1461
18.457 | -103.1512
18.4607 | -103.1551
18.4644 | -103.1593
18.4681 | -103.1647
18.4718 | -103.1677
18.4755 | -103.168
18.4792 | -103.1678
18.483 | -103.1712
18.4867 | -103.1752
18.4904 | -103.1768
18.4941 | -103.1799
18.4978 | -103.1817
18.5015 | -103.1822
18.5052 | -103.1805
18.5089 | -103.1782
18.5126 | -103.174
18.5164 | -103.1717
18.5201 | -103.1696
18.5238 | -103.1624
18.5275 | -103.1579
18.5312 | -103.158
18.5349 | -103.1538
18.5386 | -103.148
18.5423 | -103.1404
18.546 | -103.1335
18.5498 | -103.1253
18.5535 | -103.1177
18.5572 | -103.1145
18.5609 | -103.1122
18.5646 | -103.1049
18.5683 | -103.0981
18.572 | -103.092
18.5757 | -103.0861
18.5794 | -103.0806
18.5832 | -103.0741
18.5869 | -103.0677
18.5906 | -103.0587
18.5943 | -103.0503
18.598 | -103.0449
18.6017 | -103.0395
18.6054 | -103.033
18.6091 | -103.0256
18.6129 | -103.018
18.6166 | -103.0127
18.6203 | -103.0085
18.624 | -103.003
18.6277 | -102.9998
18.6314 | -102.9972
18.6351 | -102.9922
18.6388 | -102.9881
18.6425 | -102.9877
18.6463 | -102.9894
18.65 | -102.9911
18.6537 | -102.9903
18.6574 | -102.9903
18.6611 | -102.9931
18.6648 | -102.9937
18.6685 | -102.9944
18.6722 | -102.9946
18.6759 | -102.9986
18.6797 | -103.0035
18.6834 | -103.0062
18.6871 | -103.0091
18.6908 | -103.0121
18.6945 | -103.0142
18.6982 | -103.0157
18.7019 | -103.0192
18.7056 | -103.0238
18.7093 | -103.0275
18.7131 | -103.0336
18.7168 | -103.0406
18.7205 | -103.0462
18.7242 | -103.0504
18.7279 | -103.0564
18.7316 | -103.0611
18.7353 | -103.0657
18.739 | -103.073
18.7428 | -103.0812
18.7465 | -103.0893
18.7502 | -103.0949
18.7539 | -103.0988
18.7576 | -103.1035
18.7613 | -103.1082
18.765 | -103.1142
18.7687 | -103.1218
18.7724 | -103.1303
18.7762 | -103.1391
18.7799 | -103.1489
18.7836 | -103.1571
18.7873 | -103.1634
18.791 | -103.1739
18.7947 | -103.1845
18.7984 | -103.19
18.8021 | -103.1964
18.8058 | -103.2029
18.8096 | -103.2092
18.8133 | -103.213
18.817 | -103.2203
18.8207 | -103.2291
18.8244 | -103.2344
18.8281 | -103.241
18.8318 | -103.2488
18.8355 | -103.2548
18.8392 | -103.2593
18.843 | -103.2642
18.8467 | -103.2674
18.8504 | -103.2703
18.8541 | -103.2732
18.8578 | -103.277
18.8615 | -103.2797
18.8652 | -103.2809
18.8689 | -103.2795
18.8727 | -103.2811
18.8764 | -103.2821
18.8801 | -103.2837
18.8838 | -103.2838
18.8875 | -103.2839
18.8912 | -103.2834
18.8949 | -103.2825
18.8986 | -103.28
18.9023 | -103.2775
18.9061 | -103.2729
18.9098 | -103.2648
18.9135 | -103.2574
18.9172 | -103.2493
18.9209 | -103.245
18.9246 | -103.2426
18.9283 | -103.2391
18.932 | -103.2352
18.9357 | -103.2309
18.9395 | -103.2267
18.9432 | -103.2206
18.9469 | -103.2151
18.9506 | -103.212
18.9543 | -103.2082
18.958 | -103.203
18.9617 | -103.1997
18.9654 | -103.195
18.9691 | -103.189
18.9729 | -103.1827
18.9766 | -103.1784
18.9803 | -103.1731
18.984 | -103.1663
18.9877 | -103.16
18.9914 | -103.1543
18.9951 | -103.1458
18.9988 | -103.1397
19.0026 | -103.1347
19.0063 | -103.1268
19.01 | -103.119
19.0137 | -103.1119
19.0174 | -103.1074
19.0211 | -103.1017
19.0248 | -103.096
19.0285 | -103.0897
19.0322 | -103.0835
19.036 | -103.0762
19.0397 | -103.0676
19.0434 | -103.0612
19.0471 | -103.057
19.0508 | -103.053
19.0545 | -103.0502
19.0582 | -103.0449
19.0619 | -103.0399
19.0656 | -103.0352
19.0694 | -103.0304
19.0731 | -103.0254
19.0768 | -103.0236
19.0805 | -103.023
19.0842 | -103.023
19.0879 | -103.0197
19.0916 | -103.0172
19.0953 | -103.0164
19.099 | -103.0165
19.1028 | -103.0164
19.1065 | -103.0172
19.1102 | -103.0207
19.1139 | -103.0244
19.1176 | -103.0276
19.1213 | -103.031
19.125 | -103.0356
19.1287 | -103.041
19.1325 | -103.0468
19.1362 | -103.0519
19.1399 | -103.0593
19.1436 | -103.0669
19.1473 | -103.0761
19.151 | -103.0848
19.1547 | -103.0943
19.1584 | -103.1054
19.1621 | -103.1183
19.1659 | -103.1315
19.1696 | -103.1467
19.1733 | -103.1622
19.177 | -103.1772
19.1807 | -103.1935
19.1844 | -103.208
19.1881 | -103.2238
19.1918 | -103.2386
19.1955 | -103.2521
19.1993 | -103.2664
19.203 | -103.2829
19.2067 | -103.2992
19.2104 | -103.3147
19.2141 | -103.3297
19.2178 | -103.343
19.2215 | -103.3575
19.2252 | -103.372
19.2289 | -103.3854
19.2327 | -103.3963
19.2364 | -103.4101
19.2401 | -103.4226
19.2438 | -103.4345
19.2475 | -103.4435
19.2512 | -103.4521
19.2549 | -103.4618
19.2586 | -103.4705
19.2624 | -103.4781
19.2661 | -103.4856
19.2698 | -103.4918
19.2735 | -103.4911
19.2772 | -103.4948
19.2809 | -103.4993
19.2846 | -103.5033
19.2883 | -103.5059
19.292 | -103.5062
19.2958 | -103.5088
19.2995 | -103.5111
19.3032 | -103.5125
19.3069 | -103.5122
19.3106 | -103.5118
19.3143 | -103.5109
19.318 | -103.5071
19.3217 | -103.5025
19.3254 | -103.4985
19.3292 | -103.4954
19.3329 | -103.4967
19.3366 | -103.4955
19.3403 | -103.4904
19.344 | -103.4872
19.3477 | -103.487
19.3514 | -103.4831
19.3551 | -103.478
19.3588 | -103.4744
19.3626 | -103.4716
19.3663 | -103.4659
19.37 | -103.4629
19.3737 | -103.461
19.3774 | -103.4606
19.3811 | -103.4594
19.3848 | -103.4558
19.3885 | -103.4524
19.3923 | -103.452
19.396 | -103.4541
19.3997 | -103.4563
19.4034 | -103.4584
19.4071 | -103.4612
19.4108 | -103.4634
19.4145 | -103.465
19.4182 | -103.4673
19.4219 | -103.4704
19.4257 | -103.4747
19.4294 | -103.4799
19.4331 | -103.4852
19.4368 | -103.4901
19.4405 | -103.4954
19.4442 | -103.4999
19.4479 | -103.5066
19.4516 | -103.5149
19.4553 | -103.5232
19.4591 | -103.5299
19.4628 | -103.5371
19.4665 | -103.5433
19.4702 | -103.5493
19.4739 | -103.5576
19.4776 | -103.5674
19.4813 | -103.576
19.485 | -103.5824
19.4887 | -103.5895
19.4925 | -103.5961
19.4962 | -103.6039
19.4999 | -103.6131
19.5036 | -103.6214
19.5073 | -103.631
19.511 | -103.6404
19.5147 | -103.6484
19.5184 | -103.6575
19.5222 | -103.6665
19.5259 | -103.6757
19.5296 | -103.685
19.5333 | -103.6945
19.537 | -103.7045
19.5407 | -103.7116
19.5444 | -103.718
19.5481 | -103.7232
19.5518 | -103.7287
19.5556 | -103.7346
19.5593 | -103.7409
19.563 | -103.7478
19.5667 | -103.7542
19.5704 | -103.7597
19.5741 | -103.7627
19.5778 | -103.7653
19.5815 | -103.7696
19.5852 | -103.7746
19.589 | -103.777
19.5927 | -103.7786
19.5964 | -103.7812
19.6001 | -103.7828
19.6038 | -103.7837
19.6075 | -103.7834
19.6112 | -103.7814
19.6149 | -103.782
19.6186 | -103.7831
19.6224 | -103.7829
19.6261 | -103.7804
19.6298 | -103.776
19.6335 | -103.7714
19.6372 | -103.7673
19.6409 | -103.7645
19.6446 | -103.762
19.6483 | -103.7582
19.6521 | -103.7548
19.6558 | -103.7517
19.6595 | -103.7495
19.6632 | -103.7488
19.6669 | -103.7485
19.6706 | -103.7483
19.6743 | -103.7461
19.678 | -103.7444
19.6817 | -103.7446
19.6855 | -103.7449
19.6892 | -103.7474
19.6929 | -103.7508
19.6966 | -103.7525
19.7003 | -103.7559
19.704 | -103.7611
19.7077 | -103.7661
19.7114 | -103.7722
19.7151 | -103.7809
19.7189 | -103.79
19.7226 | -103.7983
19.7263 | -103.8063
19.73 | -103.8163
19.7337 | -103.8266
19.7374 | -103.8393
19.7411 | -103.8517
19.7448 | -103.8653
19.7486 | -103.8794
19.7523 | -103.8918
19.756 | -103.9036
19.7597 | -103.9174
19.7634 | -103.9335
19.7671 | -103.9493
19.7708 | -103.9642
19.7745 | -103.9764
19.7782 | -103.9886
19.782 | -104.003
19.7857 | -104.0167
19.7894 | -104.0273
19.7931 | -104.04
19.7968 | -104.0506
19.8005 | -104.0621
19.8042 | -104.0742
19.8079 | -104.0832
19.8116 | -104.0941
19.8154 | -104.1056
19.8191 | -104.1172
19.8228 | -104.128
19.8265 | -104.1383
19.8302 | -104.1494
19.8339 | -104.1587
19.8376 | -104.1676
19.8413 | -104.1747
19.845 | -104.1805
19.8488 | -104.1869
19.8525 | -104.1941
19.8562 | -104.2031
19.8599 | -104.2135
19.8636 | -104.2218
19.8673 | -104.2288
19.871 | -104.2363
19.8747 | -104.2422
19.8785 | -104.2488
19.8822 | -104.2562
19.8859 | -104.2616
19.8896 | -104.2643
19.8933 | -104.2674
19.897 | -104.2702
19.9007 | -104.2743
19.9044 | -104.2782
19.9081 | -104.2822
19.9119 | -104.2874
19.9156 | -104.2937
19.9193 | -104.2988
19.923 | -104.3029
19.9267 | -104.3073
19.9304 | -104.3092
19.9341 | -104.3101
19.9378 | -104.3109
19.9415 | -104.3115
19.9453 | -104.3121
19.949 | -104.3126
19.9527 | -104.313
19.9564 | -104.3156
19.9601 | -104.3181
19.9638 | -104.3191
19.9675 | -104.3209
19.9712 | -104.3202
19.9749 | -104.3206
19.9787 | -104.3236
19.9824 | -104.3253
19.9861 | -104.3243
19.9898 | -104.3236
19.9935 | -104.3246
19.9972 | -104.3251
20.0009 | -104.3244
20.0046 | -104.3248
20.0084 | -104.3267
20.0121 | -104.3282
20.0158 | -104.3301
20.0195 | -104.3326
20.0232 | -104.335
20.0269 | -104.3373
20.0306 | -104.3416
20.0343 | -104.3451
20.038 | -104.3462
20.0418 | -104.3491
20.0455 | -104.3523
20.0492 | -104.3551
20.0529 | -104.3579
20.0566 | -104.3617
20.0603 | -104.366
20.064 | -104.3707
20.0677 | -104.3739
20.0714 | -104.3779
20.0752 | -104.3822
20.0789 | -104.3875
20.0826 | -104.3914
20.0863 | -104.3939
20.09 | -104.3959
20.0937 | -104.3964
20.0974 | -104.3958
20.1011 | -104.3964
20.1048 | -104.3972
20.1086 | -104.399
20.1123 | -104.4013
20.116 | -104.404
20.1197 | -104.4054
20.1234 | -104.4046
20.1271 | -104.4045
20.1308 | -104.4059
20.1345 | -104.406
20.1383 | -104.4039
20.142 | -104.4035
20.1457 | -104.4015
20.1494 | -104.3992
20.1531 | -104.397
20.1568 | -104.3984
20.1605 | -104.4018
20.1642 | -104.4019
20.1679 | -104.3999
20.1717 | -104.4001
20.1754 | -104.403
20.1791 | -104.4063
20.1828 | -104.4091
20.1865 | -104.4117
20.1902 | -104.4134
20.1939 | -104.4136
20.1976 | -104.4128
20.2013 | -104.4119
20.2051 | -104.4137
20.2088 | -104.4149
20.2125 | -104.4162
20.2162 | -104.4181
20.2199 | -104.4203
20.2236 | -104.424
20.2273 | -104.4281
20.231 | -104.4295
20.2347 | -104.4315
20.2385 | -104.4331
20.2422 | -104.4333
20.2459 | -104.4345
20.2496 | -104.4381
20.2533 | -104.4388
20.257 | -104.4413
20.2607 | -104.4445
20.2644 | -104.4472
20.2682 | -104.4538
20.2719 | -104.4614
20.2756 | -104.468
20.2793 | -104.4724
20.283 | -104.4764
20.2867 | -104.4824
20.2904 | -104.4883
20.2941 | -104.4936
20.2978 | -104.5011
20.3016 | -104.5086
20.3053 | -104.5139
20.309 | -104.5218
20.3127 | -104.5289
20.3164 | -104.5343
20.3201 | -104.5423
20.3238 | -104.5528
20.3275 | -104.5618
20.3312 | -104.5721
20.335 | -104.5822
20.3387 | -104.5913
20.3424 | -104.5978
20.3461 | -104.6026
20.3498 | -104.6076
20.3535 | -104.6129
20.3572 | -104.6208
20.3609 | -104.6274
20.3646 | -104.6327
20.3684 | -104.6364
20.3721 | -104.6429
20.3758 | -104.6503
20.3795 | -104.6575
20.3832 | -104.6656
20.3869 | -104.6743
20.3906 | -104.6816
20.3943 | -104.688
20.3981 | -104.694
20.4018 | -104.6997
20.4055 | -104.7053
20.4092 | -104.7111
20.4129 | -104.7177
20.4166 | -104.7227
20.4203 | -104.7286
20.424 | -104.7358
20.4277 | -104.7449
20.4315 | -104.7499
20.4352 | -104.7569
20.4389 | -104.7662
20.4426 | -104.7741
20.4463 | -104.7777
20.45 | -104.7788
20.4537 | -104.7831
20.4574 | -104.7896
20.4611 | -104.7946
20.4649 | -104.7977
20.4686 | -104.802
20.4723 | -104.8053
20.476 | -104.8058
20.4797 | -104.8088
20.4834 | -104.8146
20.4871 | -104.8199
20.4908 | -104.8237
20.4945 | -104.8222
20.4983 | -104.8218
20.502 | -104.8231
20.5057 | -104.8238
20.5094 | -104.825
20.5131 | -104.8264
20.5168 | -104.8266
20.5205 | -104.826
20.5242 | -104.8255
20.528 | -104.8264
20.5317 | -104.8285
20.5354 | -104.8281
20.5391 | -104.8261
20.5428 | -104.8262
20.5465 | -104.8251
20.5502 | -104.823
20.5539 | -104.8206
20.5576 | -104.8141
20.5614 | -104.8089
20.5651 | -104.8049
20.5688 | -104.8022
20.5725 | -104.8003
20.5762 | -104.799
20.5799 | -104.7975
20.5836 | -104.7932
20.5873 | -104.7872
20.591 | -104.7814
20.5948 | -104.7763
20.5985 | -104.7698
20.6022 | -104.7649
20.6059 | -104.7611
20.6096 | -104.7567
20.6133 | -104.7521
20.617 | -104.7488
20.6207 | -104.7466
20.6244 | -104.7451
20.6282 | -104.7438
20.6319 | -104.7428
20.6356 | -104.7427
20.6393 | -104.7413
20.643 | -104.7412
20.6467 | -104.7406
20.6504 | -104.7386
20.6541 | -104.7382
20.6579 | -104.7385
20.6616 | -104.7394
20.6653 | -104.7378
20.669 | -104.7356
20.6727 | -104.7321
20.6764 | -104.7312
20.6801 | -104.7307
20.6838 | -104.7289
20.6875 | -104.7266
20.6913 | -104.7237
20.695 | -104.721
20.6987 | -104.7197
20.7024 | -104.7194
20.7061 | -104.7203
20.7098 | -104.7191
20.7135 | -104.7163
20.7172 | -104.7131
20.7209 | -104.7108
20.7247 | -104.7103
20.7284 | -104.71
20.7321 | -104.7089
20.7358 | -104.7074
20.7395 | -104.7051
20.7432 | -104.7053
20.7469 | -104.7055
20.7506 | -104.7025
20.7543 | -104.6982
20.7581 | -104.6967
20.7618 | -104.6933
20.7655 | -104.6893
20.7692 | -104.6881
20.7729 | -104.687
20.7766 | -104.6828
20.7803 | -104.6791
20.784 | -104.6763
20.7878 | -104.6745
20.7915 | -104.6717
20.7952 | -104.6683
20.7989 | -104.6659
20.8026 | -104.6614
20.8063 | -104.6532
20.81 | -104.6461
20.8137 | -104.6418
20.8174 | -104.6361
20.8212 | -104.6293
20.8249 | -104.6188
20.8286 | -104.6099
20.8323 | -104.6032
20.836 | -104.597
20.8397 | -104.5902
20.8434 | -104.5826
20.8471 | -104.5753
20.8508 | -104.5681
20.8546 | -104.5609
20.8583 | -104.554
20.862 | -104.5455
20.8657 | -104.5371
20.8694 | -104.5299
20.8731 | -104.5229
20.8768 | -104.5172
20.8805 | -104.5169
20.8843 | -104.5145
20.888 | -104.5131
20.8917 | -104.5134
20.8954 | -104.5128
20.8991 | -104.5118
20.9028 | -104.5112
20.9065 | -104.5129
20.9102 | -104.5149
20.9139 | -104.5181
20.9177 | -104.5196
20.9214 | -104.5241
20.9251 | -104.531
20.9288 | -104.5396
20.9325 | -104.5474
20.9362 | -104.5578
20.9399 | -104.5687
20.9436 | -104.5804
20.9473 | -104.5923
20.9511 | -104.6034
20.9548 | -104.615
20.9585 | -104.6281
20.9622 | -104.6413
20.9659 | -104.6563
20.9696 | -104.67
20.9733 | -104.6819
20.977 | -104.6961
20.9807 | -104.7109
20.9845 | -104.7275
20.9882 | -104.7452
20.9919 | -104.7607
20.9956 | -104.7736
20.9993 | -104.7864
21.003 | -104.798
21.0067 | -104.8101
21.0104 | -104.8212
21.0142 | -104.8337
21.0179 | -104.8446
21.0216 | -104.8531
21.0253 | -104.8597
21.029 | -104.8663
21.0327 | -104.875
21.0364 | -104.883
21.0401 | -104.8891
21.0438 | -104.8941
21.0476 | -104.9009
21.0513 | -104.9078
21.055 | -104.9111
21.0587 | -104.9148
21.0624 | -104.9183
21.0661 | -104.9166
21.0698 | -104.9143
21.0735 | -104.9144
21.0772 | -104.9153
21.081 | -104.9154
21.0847 | -104.9159
21.0884 | -104.9153
21.0921 | -104.9138
21.0958 | -104.9122
21.0995 | -104.9123
21.1032 | -104.9137
21.1069 | -104.9147
21.1106 | -104.9165
21.1144 | -104.9184
21.1181 | -104.9204
21.1218 | -104.9268
21.1255 | -104.9323
21.1292 | -104.936
21.1329 | -104.9378
21.1366 | -104.9422
21.1403 | -104.9477
21.1441 | -104.9543
21.1478 | -104.963
21.1515 | -104.9714
21.1552 | -104.9803
21.1589 | -104.989
21.1626 | -104.9985
21.1663 | -105.0088
21.17 | -105.0195
21.1737 | -105.031
21.1775 | -105.0426
21.1812 | -105.0525
21.1849 | -105.0638
21.1886 | -105.0755
21.1923 | -105.0868
21.196 | -105.0997
21.1997 | -105.1119
21.2034 | -105.1246
21.2071 | -105.1381
21.2109 | -105.1515
21.2146 | -105.1636
21.2183 | -105.1779
21.222 | -105.1901
21.2257 | -105.2009
21.2294 | -105.2124
21.2331 | -105.2242
21.2368 | -105.2336
21.2405 | -105.2422
21.2443 | -105.2508
21.248 | -105.2623
21.2517 | -105.2758
21.2554 | -105.2888
21.2591 | -105.2995
21.2628 | -105.3108
21.2665 | -105.3221
21.2702 | -105.3349
21.2739 | -105.3457
21.2777 | -105.3528
21.2814 | -105.3611
21.2851 | -105.3676
21.2888 | -105.374
21.2925 | -105.3813
21.2962 | -105.3892
21.2999 | -105.3972
21.3036 | -105.4064
21.3074 | -105.4159
21.3111 | -105.4235
21.3148 | -105.4291
21.3185 | -105.4342
21.3222 | -105.4411
21.3259 | -105.4479
21.3296 | -105.4522
21.3333 | -105.4557
21.337 | -105.4594
21.3408 | -105.4619
21.3445 | -105.4626
21.3482 | -105.4657
21.3519 | -105.4684
21.3556 | -105.4687
21.3593 | -105.4687
21.363 | -105.4689
21.3667 | -105.47
21.3704 | -105.4704
21.3742 | -105.4713
21.3779 | -105.4694
21.3816 | -105.4668
21.3853 | -105.4641
21.389 | -105.4611
21.3927 | -105.4569
21.3964 | -105.4521
21.4001 | -105.4479
21.4039 | -105.4434
21.4076 | -105.4374
21.4113 | -105.4295
21.415 | -105.4204
21.4187 | -105.4136
21.4224 | -105.4065
21.4261 | -105.399
21.4298 | -105.3925
21.4335 | -105.3867
21.4373 | -105.3828
21.441 | -105.3756
21.4447 | -105.3674
21.4484 | -105.3611
21.4521 | -105.3537
21.4558 | -105.3472
21.4595 | -105.3416
21.4632 | -105.3383
21.4669 | -105.335
21.4707 | -105.3296
21.4744 | -105.3259
21.4781 | -105.323
21.4818 | -105.3197
21.4855 | -105.3165
21.4892 | -105.3144
21.4929 | -105.3135
21.4966 | -105.3121
21.5003 | -105.3125
21.5041 | -105.3164
21.5078 | -105.3195
21.5115 | -105.3197
21.5152 | -105.3224
21.5189 | -105.3253
21.5226 | -105.3294
21.5263 | -105.3348
21.53 | -105.3423
21.5338 | -105.349
21.5375 | -105.3555
21.5412 | -105.3633
21.5449 | -105.3713
21.5486 | -105.3792
21.5523 | -105.3843
21.556 | -105.3896
21.5597 | -105.3966
21.5634 | -105.405
21.5672 | -105.4129
21.5709 | -105.4194
21.5746 | -105.4281
21.5783 | -105.4353
21.582 | -105.4411
21.5857 | -105.4445
21.5894 | -105.4487
21.5931 | -105.4533
21.5968 | -105.4572
21.6006 | -105.4605
21.6043 | -105.4625
21.608 | -105.4635
21.6117 | -105.4654
21.6154 | -105.4679
21.6191 | -105.4676
21.6228 | -105.4657
21.6265 | -105.4615
21.6302 | -105.4586
21.634 | -105.4571
21.6377 | -105.4534
21.6414 | -105.4486
21.6451 | -105.4434
21.6488 | -105.4401
21.6525 | -105.4337
21.6562 | -105.428
21.6599 | -105.4258
21.6637 | -105.4222
21.6674 | -105.4152
21.6711 | -105.4121
21.6748 | -105.407
21.6785 | -105.4007
21.6822 | -105.3964
21.6859 | -105.3928
21.6896 | -105.3906
21.6933 | -105.3869
21.6971 | -105.3848
21.7008 | -105.3833
21.7045 | -105.3816
21.7082 | -105.38
21.7119 | -105.3804
21.7156 | -105.3799
21.7193 | -105.3781
21.723 | -105.3791
21.7267 | -105.3801
21.7305 | -105.3832
21.7342 | -105.3858
21.7379 | -105.3883
21.7416 | -105.3899
21.7453 | -105.3912
21.749 | -105.3968
21.7527 | -105.4043
21.7564 | -105.4085
21.7601 | -105.4138
21.7639 | -105.4191
21.7676 | -105.4245
21.7713 | -105.43
21.775 | -105.4364
21.7787 | -105.4448
21.7824 | -105.4521
21.7861 | -105.4577
21.7898 | -105.4643
21.7936 | -105.4726
21.7973 | -105.4811
21.801 | -105.4893
21.8047 | -105.4998
21.8084 | -105.509
21.8121 | -105.517
21.8158 | -105.5251
21.8195 | -105.5349
21.8232 | -105.5436
21.827 | -105.5525
21.8307 | -105.5593
21.8344 | -105.5665
21.8381 | -105.5736
21.8418 | -105.5812
21.8455 | -105.5885
21.8492 | -105.5948
21.8529 | -105.6006
21.8566 | -105.6057
21.8604 | -105.6117
21.8641 | -105.6181
21.8678 | -105.6217
21.8715 | -105.6243
21.8752 | -105.6274
21.8789 | -105.631
21.8826 | -105.6333
21.8863 | -105.6354
21.89 | -105.6388
21.8938 | -105.6431
21.8975 | -105.6472
21.9012 | -105.6495
21.9049 | -105.651
21.9086 | -105.6521
21.9123 | -105.6533
21.916 | -105.6548
21.9197 | -105.6548
21.9235 | -105.6564
21.9272 | -105.6568
21.9309 | -105.6581
21.9346 | -105.663
21.9383 | -105.6647
21.942 | -105.6648
21.9457 | -105.6661
21.9494 | -105.6676
21.9531 | -105.6694
21.9569 | -105.6729
21.9606 | -105.6744
21.9643 | -105.6751
21.968 | -105.6761
21.9717 | -105.6774
21.9754 | -105.6789
21.9791 | -105.6843
21.9828 | -105.6859
21.9865 | -105.6869
21.9903 | -105.6877
21.994 | -105.6908
21.9977 | -105.6928
22.0014 | -105.6936
22.0051 | -105.6953
22.0088 | -105.6951
22.0125 | -105.6935
22.0162 | -105.6922
22.0199 | -105.6905
22.0237 | -105.6889
22.0274 | -105.689
22.0311 | -105.6897
22.0348 | -105.6862
22.0385 | -105.6831
22.0422 | -105.678
22.0459 | -105.6756
22.0496 | -105.6738
22.0534 | -105.6714
22.0571 | -105.6694
22.0608 | -105.6651
22.0645 | -105.66
22.0682 | -105.6559
22.0719 | -105.6524
22.0756 | -105.6514
22.0793 | -105.649
22.083 | -105.6464
22.0868 | -105.6437
22.0905 | -105.642
22.0942 | -105.6406
22.0979 | -105.6415
22.1016 | -105.6425
22.1053 | -105.6428
22.109 | -105.6431
22.1127 | -105.6435
22.1164 | -105.6437
22.1202 | -105.6455
22.1239 | -105.6474
22.1276 | -105.6514
22.1313 | -105.6549
22.135 | -105.6592
22.1387 | -105.6626
22.1424 | -105.6682
22.1461 | -105.6733
22.1498 | -105.6782
22.1536 | -105.6837
22.1573 | -105.6896
22.161 | -105.6969
22.1647 | -105.7051
22.1684 | -105.7132
22.1721 | -105.7201
22.1758 | -105.7267
22.1795 | -105.7351
22.1833 | -105.7448
22.187 | -105.7525
22.1907 | -105.7583
22.1944 | -105.7665
22.1981 | -105.7762
22.2018 | -105.7851
22.2055 | -105.7956
22.2092 | -105.804
22.2129 | -105.8114
22.2167 | -105.8193
22.2204 | -105.8285
22.2241 | -105.8376
22.2278 | -105.8462
22.2315 | -105.8553
22.2352 | -105.8639
22.2389 | -105.8726
22.2426 | -105.8804
22.2463 | -105.8856
22.2501 | -105.8899
22.2538 | -105.8934
22.2575 | -105.8979
22.2612 | -105.9035
22.2649 | -105.9065
22.2686 | -105.9061
22.2723 | -105.9061
22.276 | -105.9058
22.2797 | -105.9056
22.2835 | -105.902
22.2872 | -105.8972
22.2909 | -105.8953
22.2946 | -105.8921
22.2983 | -105.8889
22.302 | -105.886
22.3057 | -105.8824
22.3094 | -105.8765
22.3132 | -105.8717
22.3169 | -105.8674
22.3206 | -105.8629
22.3243 | -105.8598
22.328 | -105.8573
22.3317 | -105.8527
22.3354 | -105.8479
22.3391 | -105.8408
22.3428 | -105.8336
22.3466 | -105.8284
22.3503 | -105.825
22.354 | -105.8233
22.3577 | -105.821
22.3614 | -105.8161
22.3651 | -105.8117
22.3688 | -105.8107
22.3725 | -105.8095
22.3762 | -105.8074
22.38 | -105.8056
22.3837 | -105.8056
22.3874 | -105.8056
22.3911 | -105.8043
22.3948 | -105.8058
22.3985 | -105.8098
22.4022 | -105.8121
22.4059 | -105.8115
22.4096 | -105.8077
22.4134 | -105.7995
22.4171 | -105.7928
22.4208 | -105.791
22.4245 | -105.7911
22.4282 | -105.7898
22.4319 | -105.7839
22.4356 | -105.7708
22.4393 | -105.7569
22.4431 | -105.7499
22.4468 | -105.742
22.4505 | -105.733
22.4542 | -105.7237
22.4579 | -105.7125
22.4616 | -105.698
22.4653 | -105.6824
22.469 | -105.6607
22.4727 | -105.636
22.4765 | -105.6135
22.4802 | -105.585
22.4839 | -105.5472
22.4876 | -105.4896
22.4913 | -105.4542
22.495 | -105.4168
22.4987 | -105.3785
22.5024 | -105.3341
22.5061 | -105.2861
22.5099 | -105.2362
22.5136 | -105.1845
22.5173 | -105.1328
22.521 | -105.08
22.5247 | -105.0267
22.5284 | -104.9748
22.5321 | -104.9219
22.5358 | -104.8701
22.5395 | -104.818
22.5433 | -104.7663
22.547 | -104.7151
22.5507 | -104.6646
22.5544 | -104.6151
22.5581 | -104.5681
22.5618 | -104.5236
22.5655 | -104.4798
22.5692 | -104.4382
22.573 | -104.3989
22.5767 | -104.3609
22.5804 | -104.325
22.5841 | -104.2921
22.5878 | -104.2627
22.5915 | -104.2354
22.5952 | -104.2106
22.5989 | -104.1892
22.6026 | -104.1722
22.6064 | -104.1564
22.6101 | -104.1415
22.6138 | -104.1299
22.6175 | -104.1227
22.6212 | -104.1191
22.6249 | -104.1184
22.6286 | -104.1207
22.6323 | -104.1255
22.636 | -104.1342
22.6398 | -104.147
22.6435 | -104.1626
22.6472 | -104.1818
22.6509 | -104.2032
22.6546 | -104.2272
22.6583 | -104.2549
22.662 | -104.2859
22.6657 | -104.3188
22.6695 | -104.353
22.6732 | -104.3889
22.6769 | -104.4273
22.6806 | -104.4683
22.6843 | -104.5116
22.688 | -104.5558
22.6917 | -104.602
22.6954 | -104.6506
22.6991 | -104.6998
22.7029 | -104.7487
22.7066 | -104.7984
22.7103 | -104.8481
22.714 | -104.8977
22.7177 | -104.945
22.7214 | -104.9946
22.7251 | -105.0456
22.7288 | -105.0996
22.7325 | -105.154
22.7363 | -105.2041
22.74 | -105.2491
22.7437 | -105.2907
22.7474 | -105.3313
22.7511 | -105.3743
22.7548 | -105.4204
22.7585 | -105.4625
22.7622 | -105.4939
22.7659 | -105.5231
22.7697 | -105.55
22.7734 | -105.5757
22.7771 | -105.6003
22.7808 | -105.6249
22.7845 | -105.647
22.7882 | -105.6708
22.7919 | -105.694
22.7956 | -105.7097
22.7994 | -105.7275
22.8031 | -105.7504
22.8068 | -105.7663
22.8105 | -105.779
22.8142 | -105.7889
22.8179 | -105.7971
22.8216 | -105.8033
22.8253 | -105.8127
22.829 | -105.824
22.8328 | -105.8324
22.8365 | -105.8378
22.8402 | -105.8405
22.8439 | -105.8421
22.8476 | -105.8418
22.8513 | -105.8396
22.855 | -105.8378
22.8587 | -105.8342
22.8624 | -105.8294
22.8662 | -105.8222
22.8699 | -105.8155
22.8736 | -105.8089
22.8773 | -105.8014
22.881 | -105.7932
22.8847 | -105.7853
22.8884 | -105.7779
22.8921 | -105.7698
22.8958 | -105.7619
22.8996 | -105.7537
22.9033 | -105.7447
22.907 | -105.7355
22.9107 | -105.7266
22.9144 | -105.7184
22.9181 | -105.7096
22.9218 | -105.7003
22.9255 | -105.6909
22.9293 | -105.6819
22.933 | -105.672
22.9367 | -105.6632
22.9404 | -105.6537
22.9441 | -105.6447
22.9478 | -105.635
22.9515 | -105.6251
22.9552 | -105.6162
22.9589 | -105.6082
22.9627 | -105.6004
22.9664 | -105.5918
22.9701 | -105.5823
22.9738 | -105.5745
22.9775 | -105.5683
22.9812 | -105.5604
22.9849 | -105.5523
22.9886 | -105.5456
22.9923 | -105.5404
22.9961 | -105.5383
22.9998 | -105.537
23.0035 | -105.5356
23.0072 | -105.5342
23.0109 | -105.532
23.0146 | -105.5305
23.0183 | -105.5314
23.022 | -105.5376
23.0257 | -105.539
23.0295 | -105.5412
23.0332 | -105.5433
23.0369 | -105.547
23.0406 | -105.5502
23.0443 | -105.5543
23.048 | -105.5584
23.0517 | -105.5632
23.0554 | -105.5668
23.0592 | -105.573
23.0629 | -105.5797
23.0666 | -105.5835
23.0703 | -105.5874
23.074 | -105.5932
23.0777 | -105.5988
23.0814 | -105.6022
23.0851 | -105.6063
23.0888 | -105.61
23.0926 | -105.615
23.0963 | -105.6189
23.1 | -105.621
23.1037 | -105.623
23.1074 | -105.6238
23.1111 | -105.6262
23.1148 | -105.6294
23.1185 | -105.6317
23.1222 | -105.6342
23.126 | -105.6351
23.1297 | -105.6362
23.1334 | -105.6393
23.1371 | -105.6402
23.1408 | -105.6435
23.1445 | -105.6457
23.1482 | -105.6472
23.1519 | -105.6499
23.1556 | -105.6526
23.1594 | -105.6556
23.1631 | -105.6575
23.1668 | -105.6595
23.1705 | -105.661
23.1742 | -105.6605
23.1779 | -105.6604
23.1816 | -105.6614
23.1853 | -105.6633
23.1891 | -105.6654
23.1928 | -105.6693
23.1965 | -105.6722
23.2002 | -105.6745
23.2039 | -105.6767
23.2076 | -105.6799
23.2113 | -105.6818
23.215 | -105.6829
23.2187 | -105.6858
23.2225 | -105.6892
23.2262 | -105.6921
23.2299 | -105.695
23.2336 | -105.6985
23.2373 | -105.7014
23.241 | -105.7042
23.2447 | -105.7077
23.2484 | -105.7124
23.2521 | -105.7168
23.2559 | -105.7227
23.2596 | -105.7301
23.2633 | -105.7356
23.267 | -105.7397
23.2707 | -105.7442
23.2744 | -105.7494
23.2781 | -105.7549
23.2818 | -105.7613
23.2855 | -105.7665
23.2893 | -105.7704
23.293 | -105.7764
23.2967 | -105.7842
23.3004 | -105.7905
23.3041 | -105.7958
23.3078 | -105.8032
23.3115 | -105.8104
23.3152 | -105.8168
23.319 | -105.8229
23.3227 | -105.8253
23.3264 | -105.8269
23.3301 | -105.8294
23.3338 | -105.8319
23.3375 | -105.8363
23.3412 | -105.8397
23.3449 | -105.8418
23.3486 | -105.8444
23.3524 | -105.8481
23.3561 | -105.8497
23.3598 | -105.8508
23.3635 | -105.8532
23.3672 | -105.8554
23.3709 | -105.8575
23.3746 | -105.8602
23.3783 | -105.862
23.382 | -105.8625
23.3858 | -105.8632
23.3895 | -105.8621
23.3932 | -105.863
23.3969 | -105.8652
23.4006 | -105.8655
23.4043 | -105.8647
23.408 | -105.8649
23.4117 | -105.8657
23.4154 | -105.866
23.4192 | -105.8657
23.4229 | -105.8625
23.4266 | -105.8614
23.4303 | -105.8621
23.434 | -105.8604
23.4377 | -105.8597
23.4414 | -105.8567
23.4451 | -105.8522
23.4489 | -105.8463
23.4526 | -105.8423
23.4563 | -105.8392
23.46 | -105.8342
23.4637 | -105.8292
23.4674 | -105.8249
23.4711 | -105.8176
23.4748 | -105.8115
23.4785 | -105.8084
23.4823 | -105.8039
23.486 | -105.7983
23.4897 | -105.7937
23.4934 | -105.7902
23.4971 | -105.7864
23.5008 | -105.7827
23.5045 | -105.782
23.5082 | -105.7802
23.5119 | -105.7775
23.5157 | -105.7747
23.5194 | -105.7734
23.5231 | -105.772
23.5268 | -105.7693
23.5305 | -105.768
23.5342 | -105.7679
23.5379 | -105.7691
23.5416 | -105.7703
23.5453 | -105.7741
23.5491 | -105.7779
23.5528 | -105.7805
23.5565 | -105.7874
23.5602 | -105.791
23.5639 | -105.7953
23.5676 | -105.8011
23.5713 | -105.8079
23.575 | -105.815
23.5788 | -105.8224
23.5825 | -105.8293
23.5862 | -105.8365
23.5899 | -105.8423
23.5936 | -105.8496
23.5973 | -105.8591
23.601 | -105.8695
23.6047 | -105.879
23.6084 | -105.89
23.6122 | -105.9007
23.6159 | -105.9105
23.6196 | -105.9213
23.6233 | -105.9323
23.627 | -105.9424
23.6307 | -105.9519
23.6344 | -105.9636
23.6381 | -105.976
23.6418 | -105.9866
23.6456 | -105.9968
23.6493 | -106.0066
23.653 | -106.017
23.6567 | -106.0286
23.6604 | -106.0417
23.6641 | -106.0545
23.6678 | -106.066
23.6715 | -106.0763
23.6752 | -106.0866
23.679 | -106.0962
23.6827 | -106.1049
23.6864 | -106.1142
23.6901 | -106.1246
23.6938 | -106.1333
23.6975 | -106.1387
23.7012 | -106.1438
23.7049 | -106.1492
23.7087 | -106.1553
23.7124 | -106.1605
23.7161 | -106.1655
23.7198 | -106.1705
23.7235 | -106.1748
23.7272 | -106.1741
23.7309 | -106.1746
23.7346 | -106.1779
23.7383 | -106.18
23.7421 | -106.1806
23.7458 | -106.1815
23.7495 | -106.1829
23.7532 | -106.1838
23.7569 | -106.1846
23.7606 | -106.1843
23.7643 | -106.183
23.768 | -106.1816
23.7717 | -106.1805
23.7755 | -106.1796
23.7792 | -106.1795
23.7829 | -106.1808
23.7866 | -106.1802
23.7903 | -106.1798
23.794 | -106.1788
23.7977 | -106.1775
23.8014 | -106.1789
23.8052 | -106.1796
23.8089 | -106.1796
23.8126 | -106.1784
23.8163 | -106.178
23.82 | -106.1799
23.8237 | -106.1813
23.8274 | -106.1826
23.8311 | -106.1846
23.8348 | -106.1856
23.8386 | -106.1859
23.8423 | -106.1881
23.846 | -106.1913
23.8497 | -106.1927
23.8534 | -106.1931
23.8571 | -106.1946
23.8608 | -106.1958
23.8645 | -106.1982
23.8682 | -106.1985
23.872 | -106.1982
23.8757 | -106.1989
23.8794 | -106.1991
23.8831 | -106.1995
23.8868 | -106.2005
23.8905 | -106.2021
23.8942 | -106.2034
23.8979 | -106.2038
23.9016 | -106.2052
23.9054 | -106.2045
23.9091 | -106.2039
23.9128 | -106.203
23.9165 | -106.2015
23.9202 | -106.2017
23.9239 | -106.2021
23.9276 | -106.2039
23.9313 | -106.2043
23.9351 | -106.2029
23.9388 | -106.2019
23.9425 | -106.2028
23.9462 | -106.2032
23.9499 | -106.2028
23.9536 | -106.1998
23.9573 | -106.1966
23.961 | -106.193
23.9647 | -106.1904
23.9685 | -106.1888
23.9722 | -106.1879
23.9759 | -106.1863
23.9796 | -106.184
23.9833 | -106.183
23.987 | -106.183
23.9907 | -106.1824
23.9944 | -106.1817
23.9981 | -106.1803
24.0019 | -106.1776
24.0056 | -106.1777
24.0093 | -106.1766
24.013 | -106.1729
24.0167 | -106.1697
24.0204 | -106.1677
24.0241 | -106.1668
24.0278 | -106.1652
24.0315 | -106.1657
24.0353 | -106.1683
24.039 | -106.1698
24.0427 | -106.1707
24.0464 | -106.1706
24.0501 | -106.1712
24.0538 | -106.172
24.0575 | -106.1719
24.0612 | -106.1722
24.0649 | -106.1758
24.0687 | -106.178
24.0724 | -106.1778
24.0761 | -106.1784
24.0798 | -106.1784
24.0835 | -106.1791
24.0872 | -106.1798
24.0909 | -106.1795
24.0946 | -106.175
24.0984 | -106.1744
24.1021 | -106.174
24.1058 | -106.1744
24.1095 | -106.1734
24.1132 | -106.1697
24.1169 | -106.1679
24.1206 | -106.1662
24.1243 | -106.1647
24.128 | -106.163
24.1318 | -106.1607
24.1355 | -106.1574
24.1392 | -106.155
24.1429 | -106.1496
24.1466 | -106.1448
24.1503 | -106.1376
24.154 | -106.1291
24.1577 | -106.1221
24.1614 | -106.1148
24.1652 | -106.1079
24.1689 | -106.1002
24.1726 | -106.092
24.1763 | -106.0843
24.18 | -106.0767
24.1837 | -106.0673
24.1874 | -106.0566
24.1911 | -106.0494
24.1948 | -106.0401
24.1986 | -106.0274
24.2023 | -106.016
24.206 | -106.005
24.2097 | -105.9946
24.2134 | -105.9831
24.2171 | -105.9703
24.2208 | -105.9572
24.2245 | -105.945
24.2283 | -105.9337
24.232 | -105.9215
24.2357 | -105.9088
24.2394 | -105.8965
24.2431 | -105.884
24.2468 | -105.8728
24.2505 | -105.861
24.2542 | -105.8489
24.2579 | -105.8342
24.2617 | -105.8009
24.2654 | -105.7551
24.2691 | -105.6987
24.2728 | -105.6212
24.2765 | -105.5263
24.2802 | -105.4063
24.2839 | -105.2974
24.2876 | -105.1973
24.2913 | -105.0997
24.2951 | -105.01
24.2988 | -104.9089
24.3025 | -104.8115
24.3062 | -104.718
24.3099 | -104.6199
24.3136 | -104.5174
24.3173 | -104.4116
24.321 | -104.3021
24.3248 | -104.1889
24.3285 | -104.0732
24.3322 | -103.9549
24.3359 | -103.8336
24.3396 | -103.7111
24.3433 | -103.5875
24.347 | -103.4617
24.3507 | -103.3366
24.3544 | -103.2127
24.3582 | -103.0883
24.3619 | -102.9634
24.3656 | -102.8403
24.3693 | -102.7188
24.373 | -102.5989
24.3767 | -102.4816
24.3804 | -102.3682
24.3841 | -102.2578
24.3878 | -102.1514
24.3916 | -102.0496
24.3953 | -101.9526
24.399 | -101.8605
24.4027 | -101.7734
24.4064 | -101.6918
24.4101 | -101.6162
24.4138 | -101.547
24.4175 | -101.4835
24.4212 | -101.4267
24.425 | -101.3767
24.4287 | -101.3331
24.4324 | -101.297
24.4361 | -101.2685
24.4398 | -101.2474
24.4435 | -101.2333
24.4472 | -101.2257
24.4509 | -101.2254
24.4547 | -101.2329
24.4584 | -101.2484
24.4621 | -101.2706
24.4658 | -101.2997
24.4695 | -101.3359
24.4732 | -101.3795
24.4769 | -101.4302
24.4806 | -101.4887
24.4843 | -101.5534
24.4881 | -101.6246
24.4918 | -101.703
24.4955 | -101.7884
24.4992 | -101.8796
24.5029 | -101.9759
24.5066 | -102.0778
24.5103 | -102.1848
24.514 | -102.2967
24.5177 | -102.4136
24.5215 | -102.535
24.5252 | -102.6593
24.5289 | -102.7877
24.5326 | -102.9195
24.5363 | -103.0526
24.54 | -103.1886
24.5437 | -103.327
24.5474 | -103.4643
24.5511 | -103.5972
24.5549 | -103.7275
24.5586 | -103.8409
24.5623 | -103.9411
24.566 | -104.032
24.5697 | -104.1175
24.5734 | -104.193
24.5771 | -104.2531
24.5808 | -104.2982
24.5846 | -104.3419
24.5883 | -104.3825
24.592 | -104.414
24.5957 | -104.44
24.5994 | -104.4574
24.6031 | -104.4648
24.6068 | -104.4702
24.6105 | -104.469
24.6142 | -104.4577
24.618 | -104.4401
24.6217 | -104.4154
24.6254 | -104.3911
24.6291 | -104.2321
24.6328 | -104.1757
24.6365 | -104.0985
24.6402 | -104.009
24.6439 | -103.9059
24.6476 | -103.8022
24.6514 | -103.6894
24.6551 | -103.5685
24.6588 | -103.4472
24.6625 | -103.3271
24.6662 | -103.2094
24.6699 | -103.0944
24.6736 | -102.981
24.6773 | -102.8702
24.681 | -102.7629
24.6848 | -102.6594
24.6885 | -102.5601
24.6922 | -102.4648
24.6959 | -102.3747
24.6996 | -102.29
24.7033 | -102.2107
24.707 | -102.1364
24.7107 | -102.0682
24.7145 | -102.006
24.7182 | -101.95
24.7219 | -101.9002
24.7256 | -101.8569
24.7293 | -101.82
24.733 | -101.7899
24.7367 | -101.7664
24.7404 | -101.7489
24.7441 | -101.7375
24.7479 | -101.7327
24.7516 | -101.7346
24.7553 | -101.7429
24.759 | -101.7583
24.7627 | -101.7805
24.7664 | -101.8091
24.7701 | -101.8441
24.7738 | -101.8842
24.7775 | -101.9306
24.7813 | -101.9829
24.785 | -102.0405
24.7887 | -102.1035
24.7924 | -102.1725
24.7961 | -102.2474
24.7998 | -102.326
24.8035 | -102.4092
24.8072 | -102.4973
24.8109 | -102.5893
24.8147 | -102.6849
24.8184 | -102.7841
24.8221 | -102.8864
24.8258 | -102.9916
24.8295 | -103.099
24.8332 | -103.2079
24.8369 | -103.3187
24.8406 | -103.4305
24.8444 | -103.5433
24.8481 | -103.6563
24.8518 | -103.7689
24.8555 | -103.8817
24.8592 | -103.994
24.8629 | -104.1041
24.8666 | -104.2122
24.8703 | -104.3184
24.874 | -104.4239
24.8778 | -104.5258
24.8815 | -104.6242
24.8852 | -104.7199
24.8889 | -104.8132
24.8926 | -104.9026
24.8963 | -104.99
24.9 | -105.0773
24.9037 | -105.1627
24.9074 | -105.2605
24.9112 | -105.3613
24.9149 | -105.4571
24.9186 | -105.5403
24.9223 | -105.6168
24.926 | -105.6925
24.9297 | -105.7647
24.9334 | -105.8155
24.9371 | -105.8489
24.9408 | -105.8743
24.9446 | -105.8885
24.9483 | -105.8977
24.952 | -105.9085
24.9557 | -105.9179
24.9594 | -105.9268
24.9631 | -105.9373
24.9668 | -105.9509
24.9705 | -105.963
24.9743 | -105.9693
24.978 | -105.9731
24.9817 | -105.9757
24.9854 | -105.9744
24.9891 | -105.9728
24.9928 | -105.9717
24.9965 | -105.9698
25.0002 | -105.9699
25.0039 | -105.9708
25.0077 | -105.9707
25.0114 | -105.9712
25.0151 | -105.9707
25.0188 | -105.9689
25.0225 | -105.9663
25.0262 | -105.9663
25.0299 | -105.9646
25.0336 | -105.9619
25.0373 | -105.9615
25.0411 | -105.961
25.0448 | -105.9606
25.0485 | -105.9594
25.0522 | -105.9591
25.0559 | -105.9574
25.0596 | -105.9569
25.0633 | -105.9564
25.067 | -105.9556
25.0707 | -105.9559
25.0745 | -105.9557
25.0782 | -105.9557
25.0819 | -105.9555
25.0856 | -105.9572
25.0893 | -105.959
25.093 | -105.9615
25.0967 | -105.9638
25.1004 | -105.9682
25.1042 | -105.9727
25.1079 | -105.9754
25.1116 | -105.977
25.1153 | -105.9785
25.119 | -105.9823
25.1227 | -105.9869
25.1264 | -105.9925
25.1301 | -105.9982
25.1338 | -106.0028
25.1376 | -106.0081
25.1413 | -106.0125
25.145 | -106.0166
25.1487 | -106.0208
25.1524 | -106.0269
25.1561 | -106.0335
25.1598 | -106.0376
25.1635 | -106.0424
25.1672 | -106.0472
25.171 | -106.052
25.1747 | -106.0556
25.1784 | -106.0595
25.1821 | -106.0643
25.1858 | -106.0684
25.1895 | -106.0729
25.1932 | -106.0765
25.1969 | -106.0781
25.2006 | -106.0786
25.2044 | -106.0798
25.2081 | -106.0828
25.2118 | -106.0835
25.2155 | -106.0848
25.2192 | -106.0876
25.2229 | -106.0893
25.2266 | -106.0897
25.2303 | -106.0892
25.2341 | -106.0891
25.2378 | -106.0889
25.2415 | -106.0867
25.2452 | -106.083
25.2489 | -106.0806
25.2526 | -106.0785
25.2563 | -106.0772
25.26 | -106.0743
25.2637 | -106.0715
25.2675 | -106.0696
25.2712 | -106.0664
25.2749 | -106.0637
25.2786 | -106.0607
25.2823 | -106.0586
25.286 | -106.056
25.2897 | -106.0527
25.2934 | -106.0492
25.2971 | -106.0458
25.3009 | -106.0422
25.3046 | -106.0391
25.3083 | -106.0333
25.312 | -106.0291
25.3157 | -106.0265
25.3194 | -106.0223
25.3231 | -106.0176
25.3268 | -106.0129
25.3305 | -106.0088
25.3343 | -106.0055
25.338 | -106.0017
25.3417 | -105.9983
25.3454 | -105.9942
25.3491 | -105.9903
25.3528 | -105.9876
25.3565 | -105.9856
25.3602 | -105.9818
25.364 | -105.9798
25.3677 | -105.978
25.3714 | -105.9769
25.3751 | -105.9771
25.3788 | -105.9785
25.3825 | -105.9777
25.3862 | -105.9789
25.3899 | -105.9797
25.3936 | -105.9784
25.3974 | -105.9767
25.4011 | -105.9767
25.4048 | -105.9764
25.4085 | -105.9749
25.4122 | -105.974
25.4159 | -105.9749
25.4196 | -105.9758
25.4233 | -105.9761
25.427 | -105.9761
25.4308 | -105.9776
25.4345 | -105.9777
25.4382 | -105.9761
25.4419 | -105.9744
25.4456 | -105.9735
25.4493 | -105.9738
25.453 | -105.9748
25.4567 | -105.9743
25.4605 | -105.9739
25.4642 | -105.976
25.4679 | -105.9778
25.4716 | -105.9779
25.4753 | -105.9764
25.479 | -105.9733
25.4827 | -105.9725
25.4864 | -105.973
25.4901 | -105.9739
25.4939 | -105.9742
25.4976 | -105.9755
25.5013 | -105.9767
25.505 | -105.9786
25.5087 | -105.9795
25.5124 | -105.9795
25.5161 | -105.9805
25.5198 | -105.9815
25.5235 | -105.9829
25.5273 | -105.9852
25.531 | -105.988
25.5347 | -105.9923
25.5384 | -105.9946
25.5421 | -105.9952
25.5458 | -105.9977
25.5495 | -106.0011
25.5532 | -106.004
25.5569 | -106.0088
25.5607 | -106.0149
25.5644 | -106.0193
25.5681 | -106.022
25.5718 | -106.0251
25.5755 | -106.0281
25.5792 | -106.0319
25.5829 | -106.0375
25.5866 | -106.0433
25.5904 | -106.0489
25.5941 | -106.054
25.5978 | -106.0583
25.6015 | -106.0635
25.6052 | -106.071
25.6089 | -106.0764
25.6126 | -106.0826
25.6163 | -106.0867
25.62 | -106.0913
25.6238 | -106.0965
25.6275 | -106.0998
25.6312 | -106.1041
25.6349 | -106.1104
25.6386 | -106.116
25.6423 | -106.1201
25.646 | -106.1253
25.6497 | -106.1309
25.6534 | -106.137
25.6572 | -106.1435
25.6609 | -106.149
25.6646 | -106.153
25.6683 | -106.1577
25.672 | -106.1631
25.6757 | -106.1657
25.6794 | -106.1666
25.6831 | -106.1704
25.6868 | -106.1749
25.6906 | -106.1774
25.6943 | -106.1814
25.698 | -106.1829
25.7017 | -106.185
25.7054 | -106.1871
25.7091 | -106.1878
25.7128 | -106.1884
25.7165 | -106.1899
25.7203 | -106.1912
25.724 | -106.1935
25.7277 | -106.1935
25.7314 | -106.1944
25.7351 | -106.1948
25.7388 | -106.194
25.7425 | -106.1924
25.7462 | -106.1909
25.7499 | -106.1891
25.7537 | -106.1888
25.7574 | -106.1907
25.7611 | -106.1922
25.7648 | -106.1918
25.7685 | -106.1903
25.7722 | -106.1888
25.7759 | -106.1893
25.7796 | -106.1881
25.7833 | -106.1877
25.7871 | -106.1856
25.7908 | -106.1849
25.7945 | -106.1832
25.7982 | -106.183
25.8019 | -106.1851
25.8056 | -106.1862
25.8093 | -106.1868
25.813 | -106.1869
25.8167 | -106.1879
25.8205 | -106.1881
25.8242 | -106.1891
25.8279 | -106.1894
25.8316 | -106.1883
25.8353 | -106.1889
25.839 | -106.1921
25.8427 | -106.1963
25.8464 | -106.1985
25.8502 | -106.201
25.8539 | -106.2014
25.8576 | -106.2038
25.8613 | -106.207
25.865 | -106.2106
25.8687 | -106.214
25.8724 | -106.2178
25.8761 | -106.2228
25.8798 | -106.2273
25.8836 | -106.2326
25.8873 | -106.2379
25.891 | -106.2436
25.8947 | -106.2506
25.8984 | -106.2565
25.9021 | -106.2633
25.9058 | -106.2728
25.9095 | -106.2805
25.9132 | -106.2887
25.917 | -106.2986
25.9207 | -106.3068
25.9244 | -106.3151
25.9281 | -106.3253
25.9318 | -106.3353
25.9355 | -106.3464
25.9392 | -106.3581
25.9429 | -106.3689
25.9466 | -106.3781
25.9504 | -106.3889
25.9541 | -106.3997
25.9578 | -106.4119
25.9615 | -106.4249
25.9652 | -106.4372
25.9689 | -106.4471
25.9726 | -106.457
25.9763 | -106.4669
25.9801 | -106.4788
25.9838 | -106.4906
25.9875 | -106.5015
25.9912 | -106.5106
25.9949 | -106.5202
25.9986 | -106.5294
26.0023 | -106.5402
26.006 | -106.5513
26.0097 | -106.5605
26.0135 | -106.5684
26.0172 | -106.5743
26.0209 | -106.5811
26.0246 | -106.5905
26.0283 | -106.5999
26.032 | -106.6082
26.0357 | -106.6146
26.0394 | -106.6206
26.0431 | -106.6287
26.0469 | -106.634
26.0506 | -106.6383
26.0543 | -106.6433
26.058 | -106.6464
26.0617 | -106.6476
26.0654 | -106.6491
26.0691 | -106.6512
26.0728 | -106.6537
26.0765 | -106.6555
26.0803 | -106.6564
26.084 | -106.655
26.0877 | -106.6506
26.0914 | -106.6472
26.0951 | -106.6468
26.0988 | -106.6448
26.1025 | -106.6424
26.1062 | -106.6392
26.11 | -106.6364
26.1137 | -106.6343
26.1174 | -106.6312
26.1211 | -106.627
26.1248 | -106.6219
26.1285 | -106.6163
26.1322 | -106.6114
26.1359 | -106.6064
26.1396 | -106.601
26.1434 | -106.5943
26.1471 | -106.589
26.1508 | -106.5822
26.1545 | -106.5768
26.1582 | -106.5723
26.1619 | -106.5679
26.1656 | -106.5606
26.1693 | -106.5517
26.173 | -106.5443
26.1768 | -106.5373
26.1805 | -106.5305
26.1842 | -106.5249
26.1879 | -106.5205
26.1916 | -106.5142
26.1953 | -106.5074
26.199 | -106.5031
26.2027 | -106.4976
26.2064 | -106.4903
26.2102 | -106.4833
26.2139 | -106.4769
26.2176 | -106.4691
26.2213 | -106.4633
26.225 | -106.4557
26.2287 | -106.4475
26.2324 | -106.4394
26.2361 | -106.4321
26.2399 | -106.4253
26.2436 | -106.4197
26.2473 | -106.414
26.251 | -106.4086
26.2547 | -106.401
26.2584 | -106.3928
26.2621 | -106.3837
26.2658 | -106.3755
26.2695 | -106.3693
26.2733 | -106.3638
26.277 | -106.3564
26.2807 | -106.3502
26.2844 | -106.3424
26.2881 | -106.3353
26.2918 | -106.3307
26.2955 | -106.3267
26.2992 | -106.3232
26.3029 | -106.32
26.3067 | -106.3165
26.3104 | -106.3125
26.3141 | -106.3084
26.3178 | -106.3058
26.3215 | -106.3028
26.3252 | -106.2995
26.3289 | -106.296
26.3326 | -106.2935
26.3363 | -106.2919
26.3401 | -106.2913
26.3438 | -106.2912
26.3475 | -106.287
26.3512 | -106.2837
26.3549 | -106.2831
26.3586 | -106.2831
26.3623 | -106.2854
26.366 | -106.2861
26.3698 | -106.2853
26.3735 | -106.2845
26.3772 | -106.2836
26.3809 | -106.2829
26.3846 | -106.2839
26.3883 | -106.2868
26.392 | -106.2889
26.3957 | -106.2888
26.3994 | -106.2885
26.4032 | -106.2905
26.4069 | -106.2921
26.4106 | -106.2948
26.4143 | -106.2972
26.418 | -106.2995
26.4217 | -106.3014
26.4254 | -106.304
26.4291 | -106.3083
26.4328 | -106.3136
26.4366 | -106.3194
26.4403 | -106.3245
26.444 | -106.3287
26.4477 | -106.3334
26.4514 | -106.3388
26.4551 | -106.3434
26.4588 | -106.3499
26.4625 | -106.3552
26.4662 | -106.3593
26.47 | -106.3644
26.4737 | -106.3691
26.4774 | -106.3748
26.4811 | -106.3795
26.4848 | -106.3846
26.4885 | -106.3895
26.4922 | -106.3964
26.4959 | -106.4019
26.4997 | -106.4077
26.5034 | -106.4134
26.5071 | -106.4209
26.5108 | -106.4276
26.5145 | -106.4331
26.5182 | -106.4376
26.5219 | -106.4435
26.5256 | -106.4504
26.5293 | -106.4564
26.5331 | -106.4626
26.5368 | -106.4702
26.5405 | -106.4788
26.5442 | -106.4875
26.5479 | -106.4954
26.5516 | -106.5014
26.5553 | -106.5075
26.559 | -106.5146
26.5627 | -106.5224
26.5665 | -106.5289
26.5702 | -106.536
26.5739 | -106.5444
26.5776 | -106.5527
26.5813 | -106.5624
26.585 | -106.5715
26.5887 | -106.5808
26.5924 | -106.5903
26.5961 | -106.599
26.5999 | -106.6064
26.6036 | -106.6144
26.6073 | -106.6225
26.611 | -106.6294
26.6147 | -106.6353
26.6184 | -106.6421
26.6221 | -106.6489
26.6258 | -106.6555
26.6296 | -106.6629
26.6333 | -106.6714
26.637 | -106.679
26.6407 | -106.6853
26.6444 | -106.6897
26.6481 | -106.6926
26.6518 | -106.6984
26.6555 | -106.7044
26.6592 | -106.7096
26.663 | -106.7132
26.6667 | -106.7133
26.6704 | -106.7125
26.6741 | -106.7118
26.6778 | -106.7116
26.6815 | -106.7111
26.6852 | -106.7103
26.6889 | -106.7114
26.6926 | -106.7113
26.6964 | -106.7105
26.7001 | -106.7107
26.7038 | -106.7109
26.7075 | -106.7098
26.7112 | -106.7068
26.7149 | -106.7026
26.7186 | -106.6991
26.7223 | -106.6984
26.7261 | -106.6963
26.7298 | -106.6901
26.7335 | -106.6847
26.7372 | -106.6807
26.7409 | -106.676
26.7446 | -106.6706
26.7483 | -106.667
26.752 | -106.6641
26.7557 | -106.6594
26.7595 | -106.6527
26.7632 | -106.6459
26.7669 | -106.6393
26.7706 | -106.6361
26.7743 | -106.632
26.778 | -106.6293
26.7817 | -106.6254
26.7854 | -106.6221
26.7891 | -106.6183
26.7929 | -106.6133
26.7966 | -106.6092
26.8003 | -106.607
26.804 | -106.6047
26.8077 | -106.6008
26.8114 | -106.5984
26.8151 | -106.5962
26.8188 | -106.595
26.8225 | -106.5954
26.8263 | -106.5963
26.83 | -106.5953
26.8337 | -106.5957
26.8374 | -106.5978
26.8411 | -106.6004
26.8448 | -106.6025
26.8485 | -106.605
26.8522 | -106.6097
26.8559 | -106.6139
26.8597 | -106.6179
26.8634 | -106.6221
26.8671 | -106.6266
26.8708 | -106.631
26.8745 | -106.6362
26.8782 | -106.6412
26.8819 | -106.6466
26.8856 | -106.6519
26.8894 | -106.6581
26.8931 | -106.6657
26.8968 | -106.6731
26.9005 | -106.6798
26.9042 | -106.685
26.9079 | -106.6911
26.9116 | -106.6989
26.9153 | -106.7066
26.919 | -106.7136
26.9228 | -106.7191
26.9265 | -106.7249
26.9302 | -106.731
26.9339 | -106.737
26.9376 | -106.7417
26.9413 | -106.7464
26.945 | -106.7533
26.9487 | -106.7585
26.9524 | -106.763
26.9562 | -106.7679
26.9599 | -106.7698
26.9636 | -106.7721
26.9673 | -106.7764
26.971 | -106.7799
26.9747 | -106.7844
26.9784 | -106.7883
26.9821 | -106.7917
26.9858 | -106.7949
26.9896 | -106.7989
26.9933 | -106.8022
26.997 | -106.8045
27.0007 | -106.8051
27.0044 | -106.8059
27.0081 | -106.8073
27.0118 | -106.8073
27.0155 | -106.8069
27.0193 | -106.8073
27.023 | -106.8075
27.0267 | -106.8093
27.0304 | -106.8074
27.0341 | -106.8043
27.0378 | -106.8034
27.0415 | -106.8019
27.0452 | -106.8023
27.0489 | -106.8022
27.0527 | -106.8013
27.0564 | -106.7998
27.0601 | -106.799
27.0638 | -106.7978
27.0675 | -106.7939
27.0712 | -106.7901
27.0749 | -106.7874
27.0786 | -106.7843
27.0823 | -106.7799
27.0861 | -106.7759
27.0898 | -106.7745
27.0935 | -106.7725
27.0972 | -106.7691
27.1009 | -106.7652
27.1046 | -106.7605
27.1083 | -106.7571
27.112 | -106.7547
27.1157 | -106.7511
27.1195 | -106.7488
27.1232 | -106.7465
27.1269 | -106.7419
27.1306 | -106.7367
27.1343 | -106.7343
27.138 | -106.7323
27.1417 | -106.7307
27.1454 | -106.7287
27.1492 | -106.7245
27.1529 | -106.7207
27.1566 | -106.7167
27.1603 | -106.7141
27.164 | -106.7113
27.1677 | -106.709
27.1714 | -106.7061
27.1751 | -106.7043
27.1788 | -106.701
27.1826 | -106.6964
27.1863 | -106.6923
27.19 | -106.6875
27.1937 | -106.6825
27.1974 | -106.6799
27.2011 | -106.6771
27.2048 | -106.6746
27.2085 | -106.6723
27.2122 | -106.6702
27.216 | -106.668
27.2197 | -106.6661
27.2234 | -106.6645
27.2271 | -106.6646
27.2308 | -106.6645
27.2345 | -106.6617
27.2382 | -106.6606
27.2419 | -106.6601
27.2457 | -106.6584
27.2494 | -106.6577
27.2531 | -106.6579
27.2568 | -106.6589
27.2605 | -106.6602
27.2642 | -106.6608
27.2679 | -106.66
27.2716 | -106.6594
27.2753 | -106.6602
27.2791 | -106.6612
27.2828 | -106.6625
27.2865 | -106.6646
27.2902 | -106.6652
27.2939 | -106.6646
27.2976 | -106.6638
27.3013 | -106.6636
27.305 | -106.6641
27.3087 | -106.6649
27.3125 | -106.6657
27.3162 | -106.6664
27.3199 | -106.6665
27.3236 | -106.6672
27.3273 | -106.6688
27.331 | -106.6685
27.3347 | -106.6666
27.3384 | -106.666
27.3421 | -106.6656
27.3459 | -106.666
27.3496 | -106.666
27.3533 | -106.6648
27.357 | -106.6637
27.3607 | -106.6612
27.3644 | -106.6602
27.3681 | -106.6591
27.3718 | -106.6562
27.3756 | -106.6536
27.3793 | -106.6518
27.383 | -106.6503
27.3867 | -106.6506
27.3904 | -106.6502
27.3941 | -106.6473
27.3978 | -106.6447
27.4015 | -106.6432
27.4052 | -106.6405
27.409 | -106.6375
27.4127 | -106.6333
27.4164 | -106.6294
27.4201 | -106.627
27.4238 | -106.6229
27.4275 | -106.62
27.4312 | -106.6178
27.4349 | -106.615
27.4386 | -106.6121
27.4424 | -106.6088
27.4461 | -106.6052
27.4498 | -106.6014
27.4535 | -106.6001
27.4572 | -106.598
27.4609 | -106.5943
27.4646 | -106.5911
27.4683 | -106.5886
27.472 | -106.5845
27.4758 | -106.5792
27.4795 | -106.5756
27.4832 | -106.5741
27.4869 | -106.5719
27.4906 | -106.5687
27.4943 | -106.5654
27.498 | -106.5617
27.5017 | -106.559
27.5055 | -106.5552
27.5092 | -106.5529
27.5129 | -106.5505
27.5166 | -106.5497
27.5203 | -106.5486
27.524 | -106.5471
27.5277 | -106.5442
27.5314 | -106.5414
27.5351 | -106.539
27.5389 | -106.5378
27.5426 | -106.5368
27.5463 | -106.5352
27.55 | -106.5344
27.5537 | -106.5335
27.5574 | -106.534
27.5611 | -106.5351
27.5648 | -106.5372
27.5685 | -106.5387
27.5723 | -106.5405
27.576 | -106.5432
27.5797 | -106.5459
27.5834 | -106.5461
27.5871 | -106.5463
27.5908 | -106.5465
27.5945 | -106.5474
27.5982 | -106.5503
27.6019 | -106.5524
27.6057 | -106.5529
27.6094 | -106.5548
27.6131 | -106.5547
27.6168 | -106.5544
27.6205 | -106.5572
27.6242 | -106.5597
27.6279 | -106.5625
27.6316 | -106.5655
27.6354 | -106.5669
27.6391 | -106.568
27.6428 | -106.5705
27.6465 | -106.5731
27.6502 | -106.5752
27.6539 | -106.578
27.6576 | -106.5807
27.6613 | -106.5821
27.665 | -106.5845
27.6688 | -106.5865
27.6725 | -106.5872
27.6762 | -106.5887
27.6799 | -106.5905
27.6836 | -106.5918
27.6873 | -106.5917
27.691 | -106.594
27.6947 | -106.5951
27.6984 | -106.5946
27.7022 | -106.5959
27.7059 | -106.5983
27.7096 | -106.6002
27.7133 | -106.6011
27.717 | -106.6018
27.7207 | -106.6032
27.7244 | -106.6053
27.7281 | -106.6069
27.7318 | -106.6082
27.7356 | -106.6111
27.7393 | -106.6147
27.743 | -106.6174
27.7467 | -106.62
27.7504 | -106.6241
27.7541 | -106.6272
27.7578 | -106.6287
27.7615 | -106.6323
27.7653 | -106.6358
27.769 | -106.639
27.7727 | -106.642
27.7764 | -106.6446
27.7801 | -106.6469
27.7838 | -106.6499
27.7875 | -106.6525
27.7912 | -106.656
27.7949 | -106.6591
27.7987 | -106.6634
27.8024 | -106.6672
27.8061 | -106.6719
27.8098 | -106.6746
27.8135 | -106.6754
27.8172 | -106.6777
27.8209 | -106.679
27.8246 | -106.6806
27.8283 | -106.6807
27.8321 | -106.6831
27.8358 | -106.6868
27.8395 | -106.6883
27.8432 | -106.6904
27.8469 | -106.691
27.8506 | -106.691
27.8543 | -106.687
27.858 | -106.6731
27.8617 | -106.6641
27.8655 | -106.6609
27.8692 | -106.6577
27.8729 | -106.6527
27.8766 | -106.6455
27.8803 | -106.6388
27.884 | -106.6333
27.8877 | -106.6268
27.8914 | -106.6183
27.8952 | -106.6081
27.8989 | -106.5982
27.9026 | -106.5878
27.9063 | -106.5772
27.91 | -106.5654
27.9137 | -106.5532
27.9174 | -106.5406
27.9211 | -106.5255
27.9248 | -106.5032
27.9286 | -106.4638
27.9323 | -106.4329
27.936 | -106.4116
27.9397 | -106.3894
27.9434 | -106.3675
27.9471 | -106.3458
27.9508 | -106.3227
27.9545 | -106.2982
27.9582 | -106.2724
27.962 | -106.2472
27.9657 | -106.2211
27.9694 | -106.1939
27.9731 | -106.1654
27.9768 | -106.1362
27.9805 | -106.1067
27.9842 | -106.078
27.9879 | -106.0489
27.9916 | -106.0165
27.9954 | -105.982
27.9991 | -105.9399
28.0028 | -105.902