## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 262 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 252
Hood | 02-92B/P | 212
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1016.9 | hPa |  | PRC PRO 01
Humidity | 59.3 | %/Hr
Internal temperature of calibration bench | 23.52 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1020 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.1 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.01 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.98 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3108 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 09/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 11/07/2018
Technical Manager | Anthony HUE | Approval |  | 16/07/2017
CEO | Robert GLEMOT | Approval |  | 16/07/2017
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3217 | 57.5278
0.012209 | -35.8672 | 38.4501
0.018313 | -34.8157 | 28.4687
0.024417 | -34.364 | 22.1122
0.030521 | -34.2583 | 18.3249
0.036626 | -34.1169 | 15.2241
0.04273 | -34.0332 | 12.9101
0.048834 | -33.9808 | 11.326
0.054938 | -33.911 | 10.125
0.061043 | -33.8372 | 8.9978
0.067147 | -33.8262 | 8.1244
0.073251 | -33.8263 | 7.3715
0.079355 | -33.8433 | 6.8137
0.08546 | -33.861 | 6.3535
0.091564 | -33.8579 | 5.9752
0.097668 | -33.8244 | 5.7175
0.10377 | -33.7774 | 5.3734
0.10988 | -33.7469 | 5.0124
0.11598 | -33.738 | 4.6531
0.12209 | -33.7547 | 4.3723
0.12819 | -33.787 | 4.1374
0.13429 | -33.8183 | 3.8967
0.1404 | -33.8351 | 3.6268
0.1465 | -33.8504 | 3.3715
0.15261 | -33.8574 | 3.1346
0.15871 | -33.8539 | 2.9398
0.16482 | -33.8396 | 2.7817
0.17092 | -33.8156 | 2.7179
0.17702 | -33.7971 | 2.6192
0.18313 | -33.7802 | 2.5348
0.18923 | -33.7681 | 2.4274
0.19534 | -33.7615 | 2.3223
0.20144 | -33.7606 | 2.2077
0.20754 | -33.7634 | 2.133
0.21365 | -33.7658 | 2.0576
0.21975 | -33.77 | 1.9925
0.22586 | -33.7778 | 1.9277
0.23196 | -33.7861 | 1.856
0.23807 | -33.7944 | 1.7795
0.24417 | -33.7964 | 1.6875
0.25027 | -33.7996 | 1.6042
0.25638 | -33.7995 | 1.5132
0.26248 | -33.7977 | 1.427
0.26859 | -33.7953 | 1.3443
0.27469 | -33.7945 | 1.2582
0.2808 | -33.8002 | 1.1941
0.2869 | -33.8045 | 1.1259
0.293 | -33.8098 | 1.0687
0.29911 | -33.8159 | 1.0299
0.30521 | -33.8218 | 0.99288
0.31132 | -33.8258 | 0.97805
0.31742 | -33.83 | 0.95509
0.32353 | -33.8334 | 0.93614
0.32963 | -33.8363 | 0.92025
0.33573 | -33.8382 | 0.90639
0.34184 | -33.8407 | 0.88932
0.34794 | -33.8439 | 0.85428
0.35405 | -33.8471 | 0.82419
0.36015 | -33.8506 | 0.79081
0.36626 | -33.8529 | 0.75394
0.37236 | -33.8553 | 0.7046
0.37846 | -33.8574 | 0.64795
0.38457 | -33.8564 | 0.58948
0.39067 | -33.8569 | 0.53965
0.39678 | -33.8553 | 0.48381
0.40288 | -33.8529 | 0.43101
0.40899 | -33.8492 | 0.37943
0.41509 | -33.8448 | 0.32758
0.42119 | -33.8399 | 0.28047
0.4273 | -33.8362 | 0.2318
0.4334 | -33.8328 | 0.18233
0.43951 | -33.8298 | 0.13368
0.44561 | -33.8283 | 0.088122
0.45172 | -33.8298 | 0.050401
0.45782 | -33.8296 | 0.0084909
0.46392 | -33.8308 | -0.030022
0.47003 | -33.8324 | -0.065107
0.47613 | -33.8345 | -0.097211
0.48224 | -33.8366 | -0.12769
0.48834 | -33.8391 | -0.14919
0.49445 | -33.8411 | -0.17401
0.50055 | -33.8436 | -0.19642
0.50665 | -33.8454 | -0.21475
0.51276 | -33.8466 | -0.23453
0.51886 | -33.8478 | -0.25181
0.52497 | -33.8469 | -0.25844
0.53107 | -33.8474 | -0.26881
0.53717 | -33.8478 | -0.28001
0.54328 | -33.8476 | -0.28532
0.54938 | -33.8477 | -0.29355
0.55549 | -33.8472 | -0.3027
0.56159 | -33.8472 | -0.31399
0.5677 | -33.8468 | -0.32465
0.5738 | -33.8464 | -0.33807
0.5799 | -33.8464 | -0.35039
0.58601 | -33.8469 | -0.36708
0.59211 | -33.8479 | -0.3899
0.59822 | -33.8486 | -0.40691
0.60432 | -33.8496 | -0.42874
0.61043 | -33.8506 | -0.45027
0.61653 | -33.8518 | -0.47212
0.62263 | -33.8529 | -0.49546
0.62874 | -33.8542 | -0.51601
0.63484 | -33.8553 | -0.5374
0.64095 | -33.8566 | -0.55868
0.64705 | -33.8577 | -0.5789
0.65316 | -33.8588 | -0.59859
0.65926 | -33.8598 | -0.62119
0.66536 | -33.8593 | -0.63945
0.67147 | -33.8596 | -0.65589
0.67757 | -33.8601 | -0.67389
0.68368 | -33.8599 | -0.69382
0.68978 | -33.8601 | -0.71176
0.69589 | -33.8597 | -0.72996
0.70199 | -33.8589 | -0.7515
0.70809 | -33.8584 | -0.77151
0.7142 | -33.8579 | -0.7906
0.7203 | -33.8571 | -0.80995
0.72641 | -33.8563 | -0.83058
0.73251 | -33.8561 | -0.85111
0.73862 | -33.8561 | -0.87515
0.74472 | -33.8563 | -0.89426
0.75082 | -33.8559 | -0.91509
0.75693 | -33.856 | -0.93831
0.76303 | -33.8559 | -0.96058
0.76914 | -33.8564 | -0.98739
0.77524 | -33.8566 | -1.0121
0.78135 | -33.857 | -1.0354
0.78745 | -33.8576 | -1.06
0.79355 | -33.858 | -1.0844
0.79966 | -33.859 | -1.1101
0.80576 | -33.8601 | -1.1403
0.81187 | -33.8604 | -1.1696
0.81797 | -33.861 | -1.2001
0.82408 | -33.8616 | -1.2271
0.83018 | -33.8622 | -1.258
0.83628 | -33.863 | -1.2884
0.84239 | -33.8629 | -1.3163
0.84849 | -33.8629 | -1.3442
0.8546 | -33.8633 | -1.3717
0.8607 | -33.8634 | -1.4008
0.86681 | -33.8635 | -1.4289
0.87291 | -33.8631 | -1.4545
0.87901 | -33.8626 | -1.4768
0.88512 | -33.8624 | -1.4997
0.89122 | -33.8623 | -1.523
0.89733 | -33.8621 | -1.5422
0.90343 | -33.8618 | -1.5634
0.90953 | -33.862 | -1.5783
0.91564 | -33.8614 | -1.5943
0.92174 | -33.8611 | -1.6085
0.92785 | -33.8606 | -1.622
0.93395 | -33.8607 | -1.6338
0.94006 | -33.8607 | -1.6475
0.94616 | -33.8605 | -1.6523
0.95226 | -33.8606 | -1.6638
0.95837 | -33.8608 | -1.6755
0.96447 | -33.8612 | -1.6859
0.97058 | -33.8612 | -1.6958
0.97668 | -33.861 | -1.7039
0.98279 | -33.861 | -1.7076
0.98889 | -33.8613 | -1.7126
0.99499 | -33.8614 | -1.7212
1.0011 | -33.8617 | -1.7262
1.0072 | -33.8616 | -1.73
1.0133 | -33.8615 | -1.734
1.0194 | -33.8616 | -1.7395
1.0255 | -33.8613 | -1.7464
1.0316 | -33.8607 | -1.7506
1.0377 | -33.8602 | -1.7566
1.0438 | -33.8595 | -1.7611
1.0499 | -33.858 | -1.7676
1.056 | -33.8569 | -1.7725
1.0621 | -33.8558 | -1.7784
1.0682 | -33.8553 | -1.7833
1.0743 | -33.8542 | -1.7912
1.0805 | -33.8536 | -1.798
1.0866 | -33.852 | -1.8088
1.0927 | -33.8507 | -1.8163
1.0988 | -33.8499 | -1.8269
1.1049 | -33.8485 | -1.8344
1.111 | -33.8469 | -1.8455
1.1171 | -33.8456 | -1.8549
1.1232 | -33.8442 | -1.869
1.1293 | -33.8431 | -1.8809
1.1354 | -33.8421 | -1.8932
1.1415 | -33.841 | -1.9072
1.1476 | -33.8396 | -1.9212
1.1537 | -33.8383 | -1.9344
1.1598 | -33.8381 | -1.9526
1.1659 | -33.8376 | -1.9646
1.172 | -33.8369 | -1.9778
1.1781 | -33.8362 | -1.9918
1.1842 | -33.8356 | -2.0074
1.1903 | -33.8363 | -2.0225
1.1964 | -33.8362 | -2.037
1.2025 | -33.8361 | -2.0522
1.2086 | -33.8365 | -2.0679
1.2147 | -33.8368 | -2.0828
1.2209 | -33.8375 | -2.0963
1.227 | -33.8393 | -2.1113
1.2331 | -33.8403 | -2.1256
1.2392 | -33.8412 | -2.1413
1.2453 | -33.8424 | -2.1534
1.2514 | -33.8437 | -2.1655
1.2575 | -33.8452 | -2.1796
1.2636 | -33.8475 | -2.191
1.2697 | -33.8492 | -2.2027
1.2758 | -33.8508 | -2.2168
1.2819 | -33.8525 | -2.2303
1.288 | -33.8542 | -2.2441
1.2941 | -33.8562 | -2.2548
1.3002 | -33.8581 | -2.2631
1.3063 | -33.8598 | -2.2731
1.3124 | -33.8616 | -2.2831
1.3185 | -33.8635 | -2.2938
1.3246 | -33.8654 | -2.3046
1.3307 | -33.8674 | -2.3141
1.3368 | -33.8692 | -2.3216
1.3429 | -33.8709 | -2.3304
1.349 | -33.8723 | -2.3389
1.3551 | -33.874 | -2.3498
1.3613 | -33.8755 | -2.3577
1.3674 | -33.8768 | -2.3661
1.3735 | -33.8781 | -2.3741
1.3796 | -33.8793 | -2.3816
1.3857 | -33.8803 | -2.3886
1.3918 | -33.8813 | -2.3983
1.3979 | -33.8826 | -2.4067
1.404 | -33.8833 | -2.4167
1.4101 | -33.8842 | -2.4256
1.4162 | -33.885 | -2.4341
1.4223 | -33.8858 | -2.4444
1.4284 | -33.8866 | -2.4527
1.4345 | -33.8872 | -2.4629
1.4406 | -33.8877 | -2.474
1.4467 | -33.8883 | -2.4834
1.4528 | -33.889 | -2.4926
1.4589 | -33.8898 | -2.5031
1.465 | -33.89 | -2.5108
1.4711 | -33.8909 | -2.5198
1.4772 | -33.8914 | -2.5319
1.4833 | -33.8921 | -2.544
1.4894 | -33.8928 | -2.5554
1.4955 | -33.8935 | -2.5669
1.5016 | -33.8941 | -2.5788
1.5078 | -33.8946 | -2.5934
1.5139 | -33.8952 | -2.6059
1.52 | -33.8957 | -2.6191
1.5261 | -33.8962 | -2.6307
1.5322 | -33.8963 | -2.6433
1.5383 | -33.8966 | -2.6554
1.5444 | -33.8973 | -2.6687
1.5505 | -33.8977 | -2.6819
1.5566 | -33.8982 | -2.6947
1.5627 | -33.8988 | -2.708
1.5688 | -33.8995 | -2.7211
1.5749 | -33.8998 | -2.7334
1.581 | -33.9001 | -2.75
1.5871 | -33.9007 | -2.7656
1.5932 | -33.9011 | -2.7784
1.5993 | -33.9014 | -2.7951
1.6054 | -33.9018 | -2.81
1.6115 | -33.9021 | -2.8269
1.6176 | -33.9025 | -2.8431
1.6237 | -33.9027 | -2.8575
1.6298 | -33.903 | -2.8729
1.6359 | -33.9034 | -2.8874
1.642 | -33.9039 | -2.9018
1.6482 | -33.9042 | -2.9187
1.6543 | -33.9047 | -2.9354
1.6604 | -33.9051 | -2.95
1.6665 | -33.9052 | -2.9649
1.6726 | -33.9055 | -2.981
1.6787 | -33.906 | -2.998
1.6848 | -33.9059 | -3.0143
1.6909 | -33.9062 | -3.0305
1.697 | -33.9064 | -3.0467
1.7031 | -33.9065 | -3.0623
1.7092 | -33.9065 | -3.0772
1.7153 | -33.9067 | -3.0933
1.7214 | -33.9063 | -3.1087
1.7275 | -33.9061 | -3.1245
1.7336 | -33.9059 | -3.1406
1.7397 | -33.9059 | -3.1558
1.7458 | -33.906 | -3.1724
1.7519 | -33.9058 | -3.1885
1.758 | -33.9057 | -3.2034
1.7641 | -33.9058 | -3.2178
1.7702 | -33.9055 | -3.2336
1.7763 | -33.9056 | -3.2473
1.7824 | -33.9054 | -3.2634
1.7885 | -33.9052 | -3.2775
1.7947 | -33.9053 | -3.2924
1.8008 | -33.9052 | -3.3053
1.8069 | -33.9051 | -3.3189
1.813 | -33.9051 | -3.3331
1.8191 | -33.9048 | -3.3473
1.8252 | -33.9046 | -3.3613
1.8313 | -33.9046 | -3.3757
1.8374 | -33.9046 | -3.3892
1.8435 | -33.9046 | -3.4015
1.8496 | -33.9043 | -3.4159
1.8557 | -33.9042 | -3.4307
1.8618 | -33.9042 | -3.4425
1.8679 | -33.9042 | -3.4556
1.874 | -33.9042 | -3.4688
1.8801 | -33.904 | -3.4818
1.8862 | -33.9042 | -3.4946
1.8923 | -33.9041 | -3.5068
1.8984 | -33.9038 | -3.5176
1.9045 | -33.9037 | -3.5287
1.9106 | -33.9034 | -3.5406
1.9167 | -33.9031 | -3.5519
1.9228 | -33.9031 | -3.5629
1.9289 | -33.9027 | -3.5719
1.9351 | -33.9028 | -3.581
1.9412 | -33.9028 | -3.5914
1.9473 | -33.9029 | -3.6035
1.9534 | -33.9028 | -3.6147
1.9595 | -33.9026 | -3.6262
1.9656 | -33.9021 | -3.637
1.9717 | -33.9019 | -3.6476
1.9778 | -33.9019 | -3.6591
1.9839 | -33.9019 | -3.6706
1.99 | -33.9019 | -3.6806
1.9961 | -33.9017 | -3.6911
2.0022 | -33.9016 | -3.7011
2.0083 | -33.9013 | -3.7111
2.0144 | -33.9013 | -3.7211
2.0205 | -33.9013 | -3.7329
2.0266 | -33.9014 | -3.7446
2.0327 | -33.9014 | -3.7558
2.0388 | -33.9017 | -3.7696
2.0449 | -33.9017 | -3.78
2.051 | -33.9018 | -3.7901
2.0571 | -33.9016 | -3.8015
2.0632 | -33.9015 | -3.8121
2.0693 | -33.9015 | -3.8233
2.0754 | -33.9017 | -3.8351
2.0816 | -33.9019 | -3.8471
2.0877 | -33.9018 | -3.8583
2.0938 | -33.902 | -3.8707
2.0999 | -33.902 | -3.8823
2.106 | -33.9022 | -3.8943
2.1121 | -33.9022 | -3.9066
2.1182 | -33.9024 | -3.9167
2.1243 | -33.9025 | -3.9273
2.1304 | -33.9026 | -3.9376
2.1365 | -33.903 | -3.9492
2.1426 | -33.9031 | -3.9625
2.1487 | -33.9035 | -3.9742
2.1548 | -33.904 | -3.9873
2.1609 | -33.9045 | -3.9994
2.167 | -33.9049 | -4.012
2.1731 | -33.9053 | -4.0239
2.1792 | -33.906 | -4.0351
2.1853 | -33.9065 | -4.048
2.1914 | -33.9069 | -4.0603
2.1975 | -33.9075 | -4.0724
2.2036 | -33.9079 | -4.0856
2.2097 | -33.9083 | -4.0976
2.2158 | -33.9093 | -4.108
2.222 | -33.9098 | -4.1182
2.2281 | -33.9101 | -4.129
2.2342 | -33.9106 | -4.1421
2.2403 | -33.9111 | -4.1523
2.2464 | -33.9118 | -4.1634
2.2525 | -33.9123 | -4.1742
2.2586 | -33.9132 | -4.1842
2.2647 | -33.914 | -4.1935
2.2708 | -33.9145 | -4.2053
2.2769 | -33.9154 | -4.2168
2.283 | -33.9159 | -4.2263
2.2891 | -33.9164 | -4.2367
2.2952 | -33.917 | -4.2466
2.3013 | -33.9176 | -4.2568
2.3074 | -33.9183 | -4.2665
2.3135 | -33.9187 | -4.2779
2.3196 | -33.9193 | -4.2866
2.3257 | -33.9197 | -4.2972
2.3318 | -33.92 | -4.3063
2.3379 | -33.9205 | -4.3177
2.344 | -33.9209 | -4.3287
2.3501 | -33.9213 | -4.3395
2.3562 | -33.9216 | -4.3479
2.3623 | -33.9219 | -4.3567
2.3685 | -33.9221 | -4.3658
2.3746 | -33.9224 | -4.3764
2.3807 | -33.9229 | -4.3854
2.3868 | -33.923 | -4.3945
2.3929 | -33.9234 | -4.4039
2.399 | -33.9237 | -4.412
2.4051 | -33.9242 | -4.4216
2.4112 | -33.9245 | -4.4307
2.4173 | -33.9248 | -4.4397
2.4234 | -33.9248 | -4.4483
2.4295 | -33.9251 | -4.457
2.4356 | -33.9252 | -4.4664
2.4417 | -33.9252 | -4.4765
2.4478 | -33.9253 | -4.4862
2.4539 | -33.9256 | -4.4942
2.46 | -33.9253 | -4.504
2.4661 | -33.9253 | -4.5133
2.4722 | -33.9255 | -4.5221
2.4783 | -33.9257 | -4.5306
2.4844 | -33.9257 | -4.5393
2.4905 | -33.9256 | -4.5486
2.4966 | -33.9253 | -4.5576
2.5027 | -33.9253 | -4.5658
2.5089 | -33.9252 | -4.5744
2.515 | -33.925 | -4.5837
2.5211 | -33.9249 | -4.5921
2.5272 | -33.9246 | -4.6013
2.5333 | -33.9244 | -4.6109
2.5394 | -33.9244 | -4.62
2.5455 | -33.9242 | -4.6295
2.5516 | -33.9241 | -4.6384
2.5577 | -33.924 | -4.6484
2.5638 | -33.9238 | -4.6588
2.5699 | -33.9236 | -4.6692
2.576 | -33.9234 | -4.6796
2.5821 | -33.9232 | -4.6897
2.5882 | -33.923 | -4.6993
2.5943 | -33.9228 | -4.7092
2.6004 | -33.9227 | -4.7193
2.6065 | -33.9225 | -4.7279
2.6126 | -33.9224 | -4.7371
2.6187 | -33.9223 | -4.7459
2.6248 | -33.9221 | -4.7551
2.6309 | -33.922 | -4.7664
2.637 | -33.922 | -4.7777
2.6431 | -33.922 | -4.7873
2.6492 | -33.9218 | -4.7974
2.6554 | -33.9218 | -4.8078
2.6615 | -33.9216 | -4.8177
2.6676 | -33.9216 | -4.8283
2.6737 | -33.9217 | -4.8391
2.6798 | -33.9217 | -4.8492
2.6859 | -33.9217 | -4.8598
2.692 | -33.9216 | -4.8696
2.6981 | -33.9215 | -4.8808
2.7042 | -33.9216 | -4.8929
2.7103 | -33.9216 | -4.9031
2.7164 | -33.9218 | -4.9139
2.7225 | -33.9219 | -4.9253
2.7286 | -33.922 | -4.9365
2.7347 | -33.922 | -4.9466
2.7408 | -33.9223 | -4.956
2.7469 | -33.9222 | -4.9674
2.753 | -33.9225 | -4.9782
2.7591 | -33.9226 | -4.9902
2.7652 | -33.9226 | -5.0018
2.7713 | -33.9227 | -5.0118
2.7774 | -33.9231 | -5.0232
2.7835 | -33.9232 | -5.034
2.7896 | -33.9233 | -5.0434
2.7958 | -33.9237 | -5.0538
2.8019 | -33.9242 | -5.0637
2.808 | -33.9245 | -5.0757
2.8141 | -33.9249 | -5.0887
2.8202 | -33.9252 | -5.1009
2.8263 | -33.9257 | -5.1119
2.8324 | -33.9259 | -5.1244
2.8385 | -33.9263 | -5.1366
2.8446 | -33.9269 | -5.1484
2.8507 | -33.9272 | -5.1605
2.8568 | -33.9277 | -5.1717
2.8629 | -33.9281 | -5.1831
2.869 | -33.9286 | -5.1945
2.8751 | -33.929 | -5.2066
2.8812 | -33.9296 | -5.2199
2.8873 | -33.9299 | -5.2313
2.8934 | -33.9303 | -5.2433
2.8995 | -33.9306 | -5.2553
2.9056 | -33.9309 | -5.2662
2.9117 | -33.9314 | -5.2779
2.9178 | -33.9321 | -5.2906
2.9239 | -33.9325 | -5.304
2.93 | -33.9329 | -5.3175
2.9361 | -33.9335 | -5.329
2.9423 | -33.9341 | -5.3411
2.9484 | -33.9346 | -5.3546
2.9545 | -33.9352 | -5.3679
2.9606 | -33.9357 | -5.3813
2.9667 | -33.9363 | -5.3932
2.9728 | -33.9366 | -5.4055
2.9789 | -33.9371 | -5.4183
2.985 | -33.9375 | -5.4321
2.9911 | -33.938 | -5.4447
2.9972 | -33.9385 | -5.4575
3.0033 | -33.9391 | -5.4706
3.0094 | -33.9396 | -5.4837
3.0155 | -33.9402 | -5.4964
3.0216 | -33.9407 | -5.5098
3.0277 | -33.9411 | -5.5236
3.0338 | -33.9417 | -5.5372
3.0399 | -33.9423 | -5.551
3.046 | -33.9426 | -5.5641
3.0521 | -33.9432 | -5.5766
3.0582 | -33.9435 | -5.5905
3.0643 | -33.944 | -5.6036
3.0704 | -33.9443 | -5.617
3.0765 | -33.9446 | -5.6293
3.0827 | -33.945 | -5.6425
3.0888 | -33.9453 | -5.6554
3.0949 | -33.9456 | -5.6684
3.101 | -33.9461 | -5.6826
3.1071 | -33.9464 | -5.6967
3.1132 | -33.9468 | -5.7108
3.1193 | -33.9472 | -5.7239
3.1254 | -33.9476 | -5.7361
3.1315 | -33.948 | -5.7504
3.1376 | -33.9483 | -5.7633
3.1437 | -33.9486 | -5.777
3.1498 | -33.949 | -5.7912
3.1559 | -33.9493 | -5.8044
3.162 | -33.9493 | -5.8185
3.1681 | -33.9494 | -5.8328
3.1742 | -33.9498 | -5.8465
3.1803 | -33.9501 | -5.8598
3.1864 | -33.9505 | -5.8721
3.1925 | -33.9508 | -5.8858
3.1986 | -33.9509 | -5.8986
3.2047 | -33.951 | -5.9126
3.2108 | -33.9512 | -5.9244
3.2169 | -33.9516 | -5.9382
3.223 | -33.9516 | -5.9516
3.2292 | -33.9516 | -5.9656
3.2353 | -33.9515 | -5.9752
3.2414 | -33.9516 | -5.9882
3.2475 | -33.9516 | -5.9993
3.2536 | -33.9519 | -6.013
3.2597 | -33.9521 | -6.0243
3.2658 | -33.9522 | -6.0351
3.2719 | -33.9524 | -6.0463
3.278 | -33.9524 | -6.0585
3.2841 | -33.9526 | -6.0704
3.2902 | -33.9527 | -6.083
3.2963 | -33.9528 | -6.0947
3.3024 | -33.9527 | -6.105
3.3085 | -33.9529 | -6.1176
3.3146 | -33.9528 | -6.1286
3.3207 | -33.953 | -6.1408
3.3268 | -33.9531 | -6.1531
3.3329 | -33.9532 | -6.1629
3.339 | -33.9533 | -6.1727
3.3451 | -33.9534 | -6.1841
3.3512 | -33.9535 | -6.1937
3.3573 | -33.9536 | -6.2051
3.3634 | -33.9537 | -6.2166
3.3696 | -33.9537 | -6.2286
3.3757 | -33.9537 | -6.2385
3.3818 | -33.954 | -6.2503
3.3879 | -33.9539 | -6.2607
3.394 | -33.9538 | -6.2712
3.4001 | -33.9539 | -6.2822
3.4062 | -33.9539 | -6.2925
3.4123 | -33.9539 | -6.3023
3.4184 | -33.954 | -6.3134
3.4245 | -33.9541 | -6.324
3.4306 | -33.954 | -6.3338
3.4367 | -33.9541 | -6.3436
3.4428 | -33.9544 | -6.3529
3.4489 | -33.9544 | -6.3641
3.455 | -33.9543 | -6.3746
3.4611 | -33.9543 | -6.3858
3.4672 | -33.9543 | -6.3971
3.4733 | -33.9543 | -6.4073
3.4794 | -33.9543 | -6.4179
3.4855 | -33.9546 | -6.4275
3.4916 | -33.9547 | -6.4373
3.4977 | -33.9546 | -6.4484
3.5038 | -33.9548 | -6.4584
3.5099 | -33.9548 | -6.4681
3.5161 | -33.955 | -6.4777
3.5222 | -33.9551 | -6.4878
3.5283 | -33.955 | -6.4979
3.5344 | -33.955 | -6.5087
3.5405 | -33.9552 | -6.5187
3.5466 | -33.9553 | -6.5286
3.5527 | -33.9553 | -6.5393
3.5588 | -33.9554 | -6.5491
3.5649 | -33.9555 | -6.5594
3.571 | -33.9555 | -6.5697
3.5771 | -33.9556 | -6.5805
3.5832 | -33.9558 | -6.5902
3.5893 | -33.956 | -6.6001
3.5954 | -33.9561 | -6.6107
3.6015 | -33.9562 | -6.6211
3.6076 | -33.9564 | -6.6308
3.6137 | -33.9565 | -6.6408
3.6198 | -33.9568 | -6.6499
3.6259 | -33.9569 | -6.6607
3.632 | -33.9571 | -6.6714
3.6381 | -33.9573 | -6.6813
3.6442 | -33.9575 | -6.6918
3.6503 | -33.9577 | -6.701
3.6565 | -33.958 | -6.7127
3.6626 | -33.9581 | -6.7235
3.6687 | -33.9584 | -6.7342
3.6748 | -33.9587 | -6.7442
3.6809 | -33.959 | -6.7547
3.687 | -33.9592 | -6.766
3.6931 | -33.9596 | -6.776
3.6992 | -33.9598 | -6.7867
3.7053 | -33.96 | -6.7969
3.7114 | -33.9602 | -6.8082
3.7175 | -33.9605 | -6.8181
3.7236 | -33.9607 | -6.827
3.7297 | -33.9608 | -6.8378
3.7358 | -33.9611 | -6.8496
3.7419 | -33.9613 | -6.8612
3.748 | -33.9614 | -6.872
3.7541 | -33.9616 | -6.883
3.7602 | -33.9619 | -6.8939
3.7663 | -33.9621 | -6.9041
3.7724 | -33.9624 | -6.9152
3.7785 | -33.9627 | -6.9257
3.7846 | -33.9629 | -6.9364
3.7907 | -33.9634 | -6.9474
3.7969 | -33.9636 | -6.959
3.803 | -33.9639 | -6.9691
3.8091 | -33.9642 | -6.9794
3.8152 | -33.9645 | -6.9907
3.8213 | -33.9647 | -7.0023
3.8274 | -33.965 | -7.0129
3.8335 | -33.9654 | -7.0245
3.8396 | -33.9658 | -7.0361
3.8457 | -33.9662 | -7.0464
3.8518 | -33.9667 | -7.0573
3.8579 | -33.967 | -7.068
3.864 | -33.9674 | -7.0787
3.8701 | -33.9678 | -7.0898
3.8762 | -33.9681 | -7.101
3.8823 | -33.9683 | -7.1117
3.8884 | -33.9688 | -7.1221
3.8945 | -33.9692 | -7.1335
3.9006 | -33.9696 | -7.1451
3.9067 | -33.9699 | -7.157
3.9128 | -33.9703 | -7.1683
3.9189 | -33.9707 | -7.179
3.925 | -33.9709 | -7.1899
3.9311 | -33.9713 | -7.1999
3.9372 | -33.9717 | -7.2115
3.9434 | -33.9721 | -7.222
3.9495 | -33.9724 | -7.2337
3.9556 | -33.9729 | -7.2449
3.9617 | -33.9732 | -7.2559
3.9678 | -33.9736 | -7.2666
3.9739 | -33.974 | -7.2779
3.98 | -33.9743 | -7.2897
3.9861 | -33.9746 | -7.3002
3.9922 | -33.9748 | -7.312
3.9983 | -33.9752 | -7.3221
4.0044 | -33.9755 | -7.3318
4.0105 | -33.9759 | -7.3428
4.0166 | -33.9764 | -7.3539
4.0227 | -33.9768 | -7.3643
4.0288 | -33.977 | -7.3757
4.0349 | -33.9773 | -7.3856
4.041 | -33.9776 | -7.3966
4.0471 | -33.9779 | -7.4069
4.0532 | -33.9783 | -7.4181
4.0593 | -33.9786 | -7.4297
4.0654 | -33.9789 | -7.4406
4.0715 | -33.9792 | -7.452
4.0776 | -33.9796 | -7.4624
4.0838 | -33.9799 | -7.4738
4.0899 | -33.9803 | -7.4851
4.096 | -33.9806 | -7.4956
4.1021 | -33.9809 | -7.5064
4.1082 | -33.9812 | -7.5168
4.1143 | -33.9816 | -7.5268
4.1204 | -33.982 | -7.5376
4.1265 | -33.9824 | -7.548
4.1326 | -33.9827 | -7.5584
4.1387 | -33.9831 | -7.5696
4.1448 | -33.9835 | -7.5806
4.1509 | -33.9837 | -7.5914
4.157 | -33.984 | -7.6027
4.1631 | -33.9842 | -7.6139
4.1692 | -33.9846 | -7.6247
4.1753 | -33.985 | -7.6356
4.1814 | -33.9852 | -7.6454
4.1875 | -33.9855 | -7.6569
4.1936 | -33.9858 | -7.6683
4.1997 | -33.9861 | -7.6784
4.2058 | -33.9863 | -7.69
4.2119 | -33.9866 | -7.6998
4.218 | -33.9869 | -7.71
4.2241 | -33.9872 | -7.7196
4.2303 | -33.9876 | -7.7309
4.2364 | -33.988 | -7.7424
4.2425 | -33.9882 | -7.7529
4.2486 | -33.9885 | -7.7641
4.2547 | -33.9886 | -7.7735
4.2608 | -33.9889 | -7.7835
4.2669 | -33.9891 | -7.7939
4.273 | -33.9894 | -7.8047
4.2791 | -33.9896 | -7.8155
4.2852 | -33.9898 | -7.8258
4.2913 | -33.99 | -7.8351
4.2974 | -33.9902 | -7.8461
4.3035 | -33.9904 | -7.8569
4.3096 | -33.9906 | -7.8675
4.3157 | -33.9908 | -7.8775
4.3218 | -33.991 | -7.8875
4.3279 | -33.9913 | -7.8979
4.334 | -33.9915 | -7.9078
4.3401 | -33.9917 | -7.9178
4.3462 | -33.992 | -7.9287
4.3523 | -33.9923 | -7.9384
4.3584 | -33.9924 | -7.947
4.3645 | -33.9926 | -7.9567
4.3707 | -33.9928 | -7.966
4.3768 | -33.993 | -7.9765
4.3829 | -33.9932 | -7.9868
4.389 | -33.9934 | -7.9966
4.3951 | -33.9937 | -8.0063
4.4012 | -33.9938 | -8.0161
4.4073 | -33.9941 | -8.0263
4.4134 | -33.9943 | -8.0365
4.4195 | -33.9945 | -8.0462
4.4256 | -33.9948 | -8.0559
4.4317 | -33.995 | -8.0655
4.4378 | -33.9952 | -8.0751
4.4439 | -33.9954 | -8.0846
4.45 | -33.9955 | -8.0947
4.4561 | -33.9956 | -8.1046
4.4622 | -33.9957 | -8.1143
4.4683 | -33.9959 | -8.1236
4.4744 | -33.9961 | -8.1338
4.4805 | -33.9963 | -8.143
4.4866 | -33.9966 | -8.1523
4.4927 | -33.9967 | -8.1626
4.4988 | -33.9969 | -8.1721
4.5049 | -33.9971 | -8.1813
4.511 | -33.9974 | -8.1904
4.5172 | -33.9977 | -8.2003
4.5233 | -33.9978 | -8.211
4.5294 | -33.9981 | -8.2204
4.5355 | -33.9983 | -8.2299
4.5416 | -33.9986 | -8.2391
4.5477 | -33.9987 | -8.2485
4.5538 | -33.9989 | -8.2573
4.5599 | -33.9992 | -8.2669
4.566 | -33.9994 | -8.2767
4.5721 | -33.9996 | -8.286
4.5782 | -33.9999 | -8.2959
4.5843 | -34.0001 | -8.3056
4.5904 | -34.0004 | -8.315
4.5965 | -34.0007 | -8.3238
4.6026 | -34.001 | -8.3345
4.6087 | -34.0012 | -8.3436
4.6148 | -34.0014 | -8.3543
4.6209 | -34.0017 | -8.3629
4.627 | -34.0021 | -8.373
4.6331 | -34.0021 | -8.382
4.6392 | -34.0024 | -8.3913
4.6453 | -34.0026 | -8.4008
4.6514 | -34.0029 | -8.4104
4.6576 | -34.0029 | -8.421
4.6637 | -34.0032 | -8.4287
4.6698 | -34.0034 | -8.4374
4.6759 | -34.0035 | -8.4473
4.682 | -34.0037 | -8.4565
4.6881 | -34.0038 | -8.4655
4.6942 | -34.0041 | -8.4748
4.7003 | -34.0043 | -8.4844
4.7064 | -34.0046 | -8.4937
4.7125 | -34.0049 | -8.5033
4.7186 | -34.005 | -8.5118
4.7247 | -34.0054 | -8.5215
4.7308 | -34.0057 | -8.5309
4.7369 | -34.0058 | -8.5406
4.743 | -34.006 | -8.551
4.7491 | -34.0062 | -8.5608
4.7552 | -34.0063 | -8.5701
4.7613 | -34.0065 | -8.5798
4.7674 | -34.0068 | -8.5885
4.7735 | -34.0069 | -8.5987
4.7796 | -34.0072 | -8.6086
4.7857 | -34.0073 | -8.6184
4.7918 | -34.0075 | -8.6271
4.7979 | -34.0078 | -8.6363
4.8041 | -34.0079 | -8.6457
4.8102 | -34.0083 | -8.6548
4.8163 | -34.0085 | -8.6644
4.8224 | -34.0086 | -8.6732
4.8285 | -34.0089 | -8.6824
4.8346 | -34.0092 | -8.692
4.8407 | -34.0094 | -8.7014
4.8468 | -34.0096 | -8.7113
4.8529 | -34.0098 | -8.7211
4.859 | -34.01 | -8.7308
4.8651 | -34.0103 | -8.7402
4.8712 | -34.0105 | -8.7497
4.8773 | -34.0107 | -8.7594
4.8834 | -34.0109 | -8.7692
4.8895 | -34.0112 | -8.7794
4.8956 | -34.0114 | -8.7894
4.9017 | -34.0115 | -8.799
4.9078 | -34.0117 | -8.809
4.9139 | -34.012 | -8.8177
4.92 | -34.0122 | -8.8268
4.9261 | -34.0126 | -8.8357
4.9322 | -34.0129 | -8.845
4.9383 | -34.0132 | -8.8542
4.9445 | -34.0135 | -8.8644
4.9506 | -34.0137 | -8.8734
4.9567 | -34.014 | -8.8828
4.9628 | -34.0142 | -8.8923
4.9689 | -34.0144 | -8.9022
4.975 | -34.0146 | -8.9117
4.9811 | -34.0149 | -8.9216
4.9872 | -34.0151 | -8.9314
4.9933 | -34.0155 | -8.9412
4.9994 | -34.0158 | -8.9512
5.0055 | -34.0162 | -8.9609
5.0116 | -34.0165 | -8.9713
5.0177 | -34.0169 | -8.9815
5.0238 | -34.0172 | -8.9921
5.0299 | -34.0176 | -9.0031
5.036 | -34.0179 | -9.0128
5.0421 | -34.0181 | -9.0226
5.0482 | -34.0184 | -9.033
5.0543 | -34.0187 | -9.0424
5.0604 | -34.0191 | -9.0528
5.0665 | -34.0194 | -9.0626
5.0726 | -34.0197 | -9.0731
5.0787 | -34.02 | -9.0834
5.0848 | -34.0203 | -9.0936
5.091 | -34.0206 | -9.1035
5.0971 | -34.0209 | -9.1137
5.1032 | -34.0213 | -9.1242
5.1093 | -34.0216 | -9.1342
5.1154 | -34.0218 | -9.1443
5.1215 | -34.0221 | -9.1544
5.1276 | -34.0224 | -9.1644
5.1337 | -34.0229 | -9.1742
5.1398 | -34.0231 | -9.1841
5.1459 | -34.0235 | -9.1946
5.152 | -34.0238 | -9.2057
5.1581 | -34.024 | -9.2165
5.1642 | -34.0243 | -9.2266
5.1703 | -34.0246 | -9.2367
5.1764 | -34.0249 | -9.2467
5.1825 | -34.0252 | -9.2571
5.1886 | -34.0255 | -9.267
5.1947 | -34.0258 | -9.2779
5.2008 | -34.0262 | -9.2879
5.2069 | -34.0266 | -9.2991
5.213 | -34.0269 | -9.3088
5.2191 | -34.0273 | -9.3192
5.2252 | -34.0275 | -9.3296
5.2314 | -34.0278 | -9.3405
5.2375 | -34.0282 | -9.3505
5.2436 | -34.0284 | -9.3608
5.2497 | -34.0288 | -9.3707
5.2558 | -34.0292 | -9.3818
5.2619 | -34.0295 | -9.3928
5.268 | -34.0298 | -9.4033
5.2741 | -34.0302 | -9.4146
5.2802 | -34.0305 | -9.4255
5.2863 | -34.0309 | -9.4358
5.2924 | -34.0312 | -9.447
5.2985 | -34.0316 | -9.4585
5.3046 | -34.032 | -9.4698
5.3107 | -34.0324 | -9.4807
5.3168 | -34.0329 | -9.4912
5.3229 | -34.0332 | -9.5012
5.329 | -34.0335 | -9.5113
5.3351 | -34.034 | -9.5215
5.3412 | -34.0344 | -9.5318
5.3473 | -34.0347 | -9.5428
5.3534 | -34.0351 | -9.5537
5.3595 | -34.0354 | -9.5645
5.3656 | -34.0358 | -9.5756
5.3717 | -34.0361 | -9.5863
5.3779 | -34.0365 | -9.5972
5.384 | -34.0368 | -9.6081
5.3901 | -34.0372 | -9.6189
5.3962 | -34.0374 | -9.6297
5.4023 | -34.0378 | -9.6406
5.4084 | -34.0382 | -9.6511
5.4145 | -34.0385 | -9.6621
5.4206 | -34.0388 | -9.6735
5.4267 | -34.0392 | -9.6849
5.4328 | -34.0396 | -9.6952
5.4389 | -34.0399 | -9.7055
5.445 | -34.0402 | -9.7169
5.4511 | -34.0406 | -9.7274
5.4572 | -34.041 | -9.7387
5.4633 | -34.0413 | -9.7491
5.4694 | -34.0416 | -9.7597
5.4755 | -34.042 | -9.7707
5.4816 | -34.0423 | -9.7814
5.4877 | -34.0427 | -9.7921
5.4938 | -34.043 | -9.8026
5.4999 | -34.0433 | -9.8134
5.506 | -34.0437 | -9.8242
5.5121 | -34.0441 | -9.8354
5.5183 | -34.0444 | -9.8456
5.5244 | -34.0446 | -9.8565
5.5305 | -34.045 | -9.8674
5.5366 | -34.0453 | -9.878
5.5427 | -34.0457 | -9.8893
5.5488 | -34.046 | -9.9004
5.5549 | -34.0462 | -9.9111
5.561 | -34.0466 | -9.9225
5.5671 | -34.047 | -9.933
5.5732 | -34.0473 | -9.944
5.5793 | -34.0476 | -9.9556
5.5854 | -34.0481 | -9.9666
5.5915 | -34.0483 | -9.9777
5.5976 | -34.0486 | -9.9888
5.6037 | -34.049 | -9.9995
5.6098 | -34.0492 | -10.0101
5.6159 | -34.0496 | -10.0211
5.622 | -34.05 | -10.0317
5.6281 | -34.0503 | -10.0426
5.6342 | -34.0506 | -10.0536
5.6403 | -34.0509 | -10.0639
5.6464 | -34.0513 | -10.0748
5.6525 | -34.0516 | -10.0853
5.6586 | -34.052 | -10.0954
5.6648 | -34.0523 | -10.1068
5.6709 | -34.0526 | -10.118
5.677 | -34.0529 | -10.1295
5.6831 | -34.0533 | -10.1399
5.6892 | -34.0536 | -10.1507
5.6953 | -34.0539 | -10.1614
5.7014 | -34.0543 | -10.1728
5.7075 | -34.0545 | -10.1838
5.7136 | -34.0549 | -10.1944
5.7197 | -34.0552 | -10.2051
5.7258 | -34.0554 | -10.2158
5.7319 | -34.0558 | -10.2266
5.738 | -34.0561 | -10.2373
5.7441 | -34.0565 | -10.2481
5.7502 | -34.0569 | -10.2585
5.7563 | -34.0572 | -10.2691
5.7624 | -34.0575 | -10.2806
5.7685 | -34.0577 | -10.2912
5.7746 | -34.058 | -10.3018
5.7807 | -34.0583 | -10.3123
5.7868 | -34.0586 | -10.323
5.7929 | -34.0589 | -10.3336
5.799 | -34.0592 | -10.3443
5.8052 | -34.0595 | -10.3551
5.8113 | -34.0597 | -10.3659
5.8174 | -34.06 | -10.3768
5.8235 | -34.0603 | -10.3873
5.8296 | -34.0606 | -10.3986
5.8357 | -34.0608 | -10.4094
5.8418 | -34.0612 | -10.4203
5.8479 | -34.0614 | -10.4313
5.854 | -34.0617 | -10.4424
5.8601 | -34.062 | -10.4529
5.8662 | -34.0623 | -10.464
5.8723 | -34.0625 | -10.4741
5.8784 | -34.0628 | -10.4848
5.8845 | -34.0632 | -10.4957
5.8906 | -34.0634 | -10.5063
5.8967 | -34.0637 | -10.5166
5.9028 | -34.0641 | -10.5274
5.9089 | -34.0643 | -10.5375
5.915 | -34.0647 | -10.5482
5.9211 | -34.065 | -10.5589
5.9272 | -34.0652 | -10.5695
5.9333 | -34.0655 | -10.5795
5.9394 | -34.0659 | -10.5899
5.9455 | -34.0662 | -10.6002
5.9517 | -34.0666 | -10.6109
5.9578 | -34.0669 | -10.6218
5.9639 | -34.0672 | -10.6321
5.97 | -34.0675 | -10.6429
5.9761 | -34.0678 | -10.6532
5.9822 | -34.0682 | -10.664
5.9883 | -34.0685 | -10.6744
5.9944 | -34.0688 | -10.685
6.0005 | -34.0692 | -10.6959
6.0066 | -34.0695 | -10.7061
6.0127 | -34.0698 | -10.716
6.0188 | -34.0701 | -10.727
6.0249 | -34.0704 | -10.7375
6.031 | -34.0707 | -10.7478
6.0371 | -34.0711 | -10.7587
6.0432 | -34.0714 | -10.7692
6.0493 | -34.0718 | -10.7796
6.0554 | -34.0722 | -10.7901
6.0615 | -34.0725 | -10.8011
6.0676 | -34.0728 | -10.8123
6.0737 | -34.0731 | -10.823
6.0798 | -34.0734 | -10.8334
6.0859 | -34.0737 | -10.8439
6.0921 | -34.0741 | -10.8542
6.0982 | -34.0744 | -10.865
6.1043 | -34.0748 | -10.8754
6.1104 | -34.0752 | -10.8864
6.1165 | -34.0755 | -10.8962
6.1226 | -34.0758 | -10.9062
6.1287 | -34.0762 | -10.9168
6.1348 | -34.0766 | -10.9274
6.1409 | -34.0769 | -10.9378
6.147 | -34.0773 | -10.9482
6.1531 | -34.0776 | -10.9586
6.1592 | -34.0779 | -10.9691
6.1653 | -34.0782 | -10.9794
6.1714 | -34.0786 | -10.9899
6.1775 | -34.0789 | -11.0007
6.1836 | -34.0792 | -11.011
6.1897 | -34.0796 | -11.0215
6.1958 | -34.0798 | -11.0322
6.2019 | -34.0802 | -11.0429
6.208 | -34.0806 | -11.0538
6.2141 | -34.0809 | -11.0645
6.2202 | -34.0813 | -11.0754
6.2263 | -34.0817 | -11.0858
6.2325 | -34.082 | -11.0958
6.2386 | -34.0824 | -11.1059
6.2447 | -34.0827 | -11.1167
6.2508 | -34.083 | -11.127
6.2569 | -34.0834 | -11.1378
6.263 | -34.0838 | -11.1479
6.2691 | -34.0841 | -11.1583
6.2752 | -34.0845 | -11.1686
6.2813 | -34.0848 | -11.1791
6.2874 | -34.0851 | -11.1895
6.2935 | -34.0855 | -11.1998
6.2996 | -34.0858 | -11.2103
6.3057 | -34.0862 | -11.2207
6.3118 | -34.0866 | -11.2311
6.3179 | -34.0869 | -11.2415
6.324 | -34.0873 | -11.2522
6.3301 | -34.0877 | -11.2624
6.3362 | -34.088 | -11.2729
6.3423 | -34.0885 | -11.2834
6.3484 | -34.0889 | -11.2941
6.3545 | -34.0892 | -11.3039
6.3606 | -34.0896 | -11.3141
6.3667 | -34.09 | -11.3242
6.3728 | -34.0904 | -11.3347
6.379 | -34.0907 | -11.3451
6.3851 | -34.0911 | -11.3556
6.3912 | -34.0915 | -11.3657
6.3973 | -34.0919 | -11.3759
6.4034 | -34.0922 | -11.3862
6.4095 | -34.0925 | -11.3968
6.4156 | -34.0929 | -11.4069
6.4217 | -34.0932 | -11.4172
6.4278 | -34.0936 | -11.4275
6.4339 | -34.094 | -11.4378
6.44 | -34.0944 | -11.4479
6.4461 | -34.0947 | -11.4587
6.4522 | -34.0951 | -11.4685
6.4583 | -34.0954 | -11.4789
6.4644 | -34.0957 | -11.4896
6.4705 | -34.0961 | -11.5001
6.4766 | -34.0965 | -11.5109
6.4827 | -34.0968 | -11.5217
6.4888 | -34.0971 | -11.5319
6.4949 | -34.0975 | -11.5419
6.501 | -34.0979 | -11.5525
6.5071 | -34.0983 | -11.5631
6.5132 | -34.0987 | -11.5733
6.5194 | -34.099 | -11.5836
6.5255 | -34.0994 | -11.5935
6.5316 | -34.0997 | -11.6041
6.5377 | -34.1001 | -11.6147
6.5438 | -34.1005 | -11.6249
6.5499 | -34.1009 | -11.6351
6.556 | -34.1012 | -11.6453
6.5621 | -34.1015 | -11.6558
6.5682 | -34.1019 | -11.666
6.5743 | -34.1022 | -11.6771
6.5804 | -34.1025 | -11.6875
6.5865 | -34.1028 | -11.6981
6.5926 | -34.1033 | -11.7084
6.5987 | -34.1036 | -11.7184
6.6048 | -34.104 | -11.7294
6.6109 | -34.1044 | -11.7397
6.617 | -34.1047 | -11.7509
6.6231 | -34.105 | -11.7615
6.6292 | -34.1054 | -11.7717
6.6353 | -34.1058 | -11.7831
6.6414 | -34.1062 | -11.7938
6.6475 | -34.1066 | -11.8043
6.6536 | -34.1069 | -11.8143
6.6597 | -34.1074 | -11.8245
6.6659 | -34.1077 | -11.8347
6.672 | -34.1081 | -11.8451
6.6781 | -34.1085 | -11.8553
6.6842 | -34.1088 | -11.8658
6.6903 | -34.1092 | -11.8761
6.6964 | -34.1096 | -11.8866
6.7025 | -34.1099 | -11.8969
6.7086 | -34.1103 | -11.9071
6.7147 | -34.1107 | -11.9176
6.7208 | -34.111 | -11.9285
6.7269 | -34.1113 | -11.9385
6.733 | -34.1116 | -11.9481
6.7391 | -34.1121 | -11.9587
6.7452 | -34.1125 | -11.9685
6.7513 | -34.1129 | -11.9785
6.7574 | -34.1133 | -11.9891
6.7635 | -34.1137 | -11.999
6.7696 | -34.1141 | -12.0094
6.7757 | -34.1145 | -12.02
6.7818 | -34.1149 | -12.0305
6.7879 | -34.1152 | -12.0409
6.794 | -34.1156 | -12.051
6.8001 | -34.1159 | -12.0612
6.8063 | -34.1163 | -12.0717
6.8124 | -34.1167 | -12.0821
6.8185 | -34.117 | -12.0925
6.8246 | -34.1174 | -12.1028
6.8307 | -34.1177 | -12.113
6.8368 | -34.1181 | -12.1236
6.8429 | -34.1185 | -12.1339
6.849 | -34.1188 | -12.1442
6.8551 | -34.1192 | -12.1546
6.8612 | -34.1196 | -12.1647
6.8673 | -34.12 | -12.1748
6.8734 | -34.1204 | -12.1859
6.8795 | -34.1208 | -12.1964
6.8856 | -34.1211 | -12.2063
6.8917 | -34.1215 | -12.216
6.8978 | -34.1218 | -12.2263
6.9039 | -34.1222 | -12.2364
6.91 | -34.1226 | -12.2468
6.9161 | -34.123 | -12.2569
6.9222 | -34.1233 | -12.2672
6.9283 | -34.1236 | -12.2772
6.9344 | -34.124 | -12.2873
6.9405 | -34.1244 | -12.2974
6.9466 | -34.1247 | -12.3079
6.9528 | -34.1251 | -12.3182
6.9589 | -34.1255 | -12.3283
6.965 | -34.1258 | -12.3386
6.9711 | -34.1262 | -12.3489
6.9772 | -34.1267 | -12.3589
6.9833 | -34.1269 | -12.3694
6.9894 | -34.1273 | -12.38
6.9955 | -34.1277 | -12.3902
7.0016 | -34.128 | -12.4009
7.0077 | -34.1284 | -12.4114
7.0138 | -34.1288 | -12.422
7.0199 | -34.1291 | -12.432
7.026 | -34.1294 | -12.4422
7.0321 | -34.1297 | -12.4524
7.0382 | -34.13 | -12.4624
7.0443 | -34.1305 | -12.473
7.0504 | -34.1308 | -12.4835
7.0565 | -34.1311 | -12.494
7.0626 | -34.1315 | -12.5041
7.0687 | -34.1318 | -12.5139
7.0748 | -34.1322 | -12.524
7.0809 | -34.1325 | -12.5343
7.087 | -34.1329 | -12.5447
7.0932 | -34.1332 | -12.5548
7.0993 | -34.1335 | -12.5651
7.1054 | -34.1339 | -12.5751
7.1115 | -34.1342 | -12.5852
7.1176 | -34.1345 | -12.5952
7.1237 | -34.1348 | -12.6054
7.1298 | -34.1351 | -12.6152
7.1359 | -34.1354 | -12.6251
7.142 | -34.1357 | -12.6353
7.1481 | -34.136 | -12.6452
7.1542 | -34.1363 | -12.6553
7.1603 | -34.1367 | -12.6653
7.1664 | -34.1371 | -12.6756
7.1725 | -34.1374 | -12.6855
7.1786 | -34.1378 | -12.6957
7.1847 | -34.1381 | -12.7057
7.1908 | -34.1384 | -12.716
7.1969 | -34.1388 | -12.7261
7.203 | -34.1391 | -12.7362
7.2091 | -34.1394 | -12.7461
7.2152 | -34.1397 | -12.7562
7.2213 | -34.1401 | -12.7661
7.2274 | -34.1404 | -12.7762
7.2335 | -34.1407 | -12.7863
7.2397 | -34.1411 | -12.7964
7.2458 | -34.1414 | -12.806
7.2519 | -34.1417 | -12.8161
7.258 | -34.1421 | -12.8263
7.2641 | -34.1424 | -12.8363
7.2702 | -34.1427 | -12.8467
7.2763 | -34.1431 | -12.8566
7.2824 | -34.1435 | -12.8666
7.2885 | -34.1439 | -12.8765
7.2946 | -34.1442 | -12.8868
7.3007 | -34.1445 | -12.8969
7.3068 | -34.1448 | -12.9072
7.3129 | -34.1451 | -12.917
7.319 | -34.1455 | -12.9271
7.3251 | -34.1458 | -12.9368
7.3312 | -34.1461 | -12.9466
7.3373 | -34.1465 | -12.9563
7.3434 | -34.1469 | -12.966
7.3495 | -34.1473 | -12.9756
7.3556 | -34.1476 | -12.986
7.3617 | -34.1479 | -12.9964
7.3678 | -34.1483 | -13.0065
7.3739 | -34.1486 | -13.0161
7.3801 | -34.1489 | -13.0261
7.3862 | -34.1492 | -13.036
7.3923 | -34.1496 | -13.0459
7.3984 | -34.1499 | -13.0558
7.4045 | -34.1502 | -13.0656
7.4106 | -34.1506 | -13.0754
7.4167 | -34.1509 | -13.0856
7.4228 | -34.1513 | -13.0955
7.4289 | -34.1516 | -13.1052
7.435 | -34.152 | -13.1149
7.4411 | -34.1523 | -13.1245
7.4472 | -34.1526 | -13.1346
7.4533 | -34.153 | -13.1445
7.4594 | -34.1533 | -13.1548
7.4655 | -34.1537 | -13.1646
7.4716 | -34.154 | -13.1748
7.4777 | -34.1543 | -13.1842
7.4838 | -34.1547 | -13.1943
7.4899 | -34.1551 | -13.2042
7.496 | -34.1554 | -13.2142
7.5021 | -34.1558 | -13.2239
7.5082 | -34.1561 | -13.2339
7.5143 | -34.1564 | -13.2437
7.5204 | -34.1567 | -13.2534
7.5266 | -34.157 | -13.2633
7.5327 | -34.1574 | -13.2731
7.5388 | -34.1577 | -13.2829
7.5449 | -34.158 | -13.293
7.551 | -34.1583 | -13.3031
7.5571 | -34.1587 | -13.3129
7.5632 | -34.159 | -13.3227
7.5693 | -34.1594 | -13.3328
7.5754 | -34.1597 | -13.3427
7.5815 | -34.1601 | -13.3522
7.5876 | -34.1604 | -13.362
7.5937 | -34.1607 | -13.3718
7.5998 | -34.161 | -13.3817
7.6059 | -34.1613 | -13.3916
7.612 | -34.1617 | -13.4014
7.6181 | -34.162 | -13.4112
7.6242 | -34.1623 | -13.4211
7.6303 | -34.1627 | -13.4308
7.6364 | -34.163 | -13.4409
7.6425 | -34.1634 | -13.4511
7.6486 | -34.1637 | -13.461
7.6547 | -34.1641 | -13.4709
7.6608 | -34.1644 | -13.4808
7.667 | -34.1647 | -13.4901
7.6731 | -34.165 | -13.5001
7.6792 | -34.1653 | -13.5101
7.6853 | -34.1656 | -13.5198
7.6914 | -34.1659 | -13.5298
7.6975 | -34.1662 | -13.5399
7.7036 | -34.1665 | -13.5499
7.7097 | -34.1669 | -13.5596
7.7158 | -34.1672 | -13.5694
7.7219 | -34.1675 | -13.5791
7.728 | -34.1678 | -13.5891
7.7341 | -34.1681 | -13.599
7.7402 | -34.1685 | -13.6088
7.7463 | -34.1688 | -13.6187
7.7524 | -34.169 | -13.6288
7.7585 | -34.1694 | -13.6388
7.7646 | -34.1697 | -13.6484
7.7707 | -34.1699 | -13.6581
7.7768 | -34.1702 | -13.6678
7.7829 | -34.1704 | -13.6778
7.789 | -34.1708 | -13.6874
7.7951 | -34.1711 | -13.6976
7.8012 | -34.1714 | -13.7078
7.8073 | -34.1717 | -13.7176
7.8135 | -34.172 | -13.7275
7.8196 | -34.1723 | -13.7373
7.8257 | -34.1727 | -13.7476
7.8318 | -34.173 | -13.7578
7.8379 | -34.1733 | -13.7672
7.844 | -34.1736 | -13.7767
7.8501 | -34.1738 | -13.7868
7.8562 | -34.1741 | -13.7973
7.8623 | -34.1745 | -13.8072
7.8684 | -34.1747 | -13.8168
7.8745 | -34.175 | -13.8265
7.8806 | -34.1753 | -13.8359
7.8867 | -34.1756 | -13.8456
7.8928 | -34.176 | -13.8553
7.8989 | -34.1763 | -13.8647
7.905 | -34.1767 | -13.875
7.9111 | -34.177 | -13.8849
7.9172 | -34.1773 | -13.8945
7.9233 | -34.1775 | -13.9044
7.9294 | -34.1779 | -13.9144
7.9355 | -34.1782 | -13.9238
7.9416 | -34.1784 | -13.9335
7.9477 | -34.1787 | -13.9429
7.9539 | -34.1791 | -13.9523
7.96 | -34.1794 | -13.9624
7.9661 | -34.1797 | -13.9723
7.9722 | -34.18 | -13.9822
7.9783 | -34.1803 | -13.992
7.9844 | -34.1807 | -14.0023
7.9905 | -34.1809 | -14.012
7.9966 | -34.1812 | -14.0218
8.0027 | -34.1815 | -14.0319
8.0088 | -34.1818 | -14.0414
8.0149 | -34.1821 | -14.051
8.021 | -34.1824 | -14.0608
8.0271 | -34.1827 | -14.0711
8.0332 | -34.183 | -14.0809
8.0393 | -34.1833 | -14.0909
8.0454 | -34.1836 | -14.1009
8.0515 | -34.1839 | -14.1107
8.0576 | -34.1842 | -14.1208
8.0637 | -34.1845 | -14.1301
8.0698 | -34.1848 | -14.1395
8.0759 | -34.1851 | -14.149
8.082 | -34.1854 | -14.159
8.0881 | -34.1857 | -14.1687
8.0942 | -34.186 | -14.179
8.1004 | -34.1863 | -14.1891
8.1065 | -34.1866 | -14.1989
8.1126 | -34.1869 | -14.2086
8.1187 | -34.1873 | -14.2186
8.1248 | -34.1876 | -14.2286
8.1309 | -34.1879 | -14.2384
8.137 | -34.1881 | -14.2481
8.1431 | -34.1885 | -14.2581
8.1492 | -34.1888 | -14.268
8.1553 | -34.1891 | -14.2779
8.1614 | -34.1894 | -14.288
8.1675 | -34.1897 | -14.2976
8.1736 | -34.19 | -14.3078
8.1797 | -34.1902 | -14.3178
8.1858 | -34.1905 | -14.3275
8.1919 | -34.1908 | -14.3373
8.198 | -34.1911 | -14.3468
8.2041 | -34.1914 | -14.3569
8.2102 | -34.1918 | -14.3668
8.2163 | -34.1921 | -14.3765
8.2224 | -34.1924 | -14.3866
8.2285 | -34.1927 | -14.3962
8.2346 | -34.193 | -14.406
8.2408 | -34.1933 | -14.4159
8.2469 | -34.1936 | -14.4257
8.253 | -34.1939 | -14.4354
8.2591 | -34.1942 | -14.445
8.2652 | -34.1946 | -14.4547
8.2713 | -34.1949 | -14.4646
8.2774 | -34.1952 | -14.4744
8.2835 | -34.1955 | -14.4847
8.2896 | -34.1958 | -14.494
8.2957 | -34.1961 | -14.5039
8.3018 | -34.1964 | -14.5137
8.3079 | -34.1968 | -14.5235
8.314 | -34.1971 | -14.5331
8.3201 | -34.1974 | -14.5429
8.3262 | -34.1977 | -14.5527
8.3323 | -34.198 | -14.5624
8.3384 | -34.1984 | -14.5726
8.3445 | -34.1987 | -14.582
8.3506 | -34.199 | -14.5921
8.3567 | -34.1993 | -14.6021
8.3628 | -34.1996 | -14.6119
8.3689 | -34.1999 | -14.6217
8.375 | -34.2002 | -14.6316
8.3812 | -34.2005 | -14.6414
8.3873 | -34.2008 | -14.6513
8.3934 | -34.2011 | -14.6611
8.3995 | -34.2014 | -14.671
8.4056 | -34.2017 | -14.6809
8.4117 | -34.202 | -14.691
8.4178 | -34.2023 | -14.7004
8.4239 | -34.2027 | -14.71
8.43 | -34.203 | -14.7199
8.4361 | -34.2033 | -14.73
8.4422 | -34.2036 | -14.7398
8.4483 | -34.2039 | -14.7496
8.4544 | -34.2042 | -14.7597
8.4605 | -34.2045 | -14.7694
8.4666 | -34.2048 | -14.7794
8.4727 | -34.205 | -14.7891
8.4788 | -34.2054 | -14.7988
8.4849 | -34.2057 | -14.809
8.491 | -34.206 | -14.8186
8.4971 | -34.2063 | -14.8285
8.5032 | -34.2066 | -14.8384
8.5093 | -34.2069 | -14.8482
8.5154 | -34.2072 | -14.8578
8.5215 | -34.2075 | -14.8677
8.5277 | -34.2078 | -14.8776
8.5338 | -34.2081 | -14.8872
8.5399 | -34.2084 | -14.8973
8.546 | -34.2087 | -14.9068
8.5521 | -34.209 | -14.9168
8.5582 | -34.2093 | -14.9267
8.5643 | -34.2096 | -14.9363
8.5704 | -34.2099 | -14.9458
8.5765 | -34.2103 | -14.9556
8.5826 | -34.2107 | -14.9654
8.5887 | -34.2109 | -14.9751
8.5948 | -34.2112 | -14.9845
8.6009 | -34.2116 | -14.9943
8.607 | -34.2118 | -15.0043
8.6131 | -34.2121 | -15.0144
8.6192 | -34.2124 | -15.0239
8.6253 | -34.2128 | -15.0338
8.6314 | -34.2131 | -15.0435
8.6375 | -34.2134 | -15.0535
8.6436 | -34.2138 | -15.0634
8.6497 | -34.2141 | -15.0733
8.6558 | -34.2144 | -15.0828
8.6619 | -34.2147 | -15.0928
8.6681 | -34.215 | -15.1025
8.6742 | -34.2153 | -15.1126
8.6803 | -34.2156 | -15.1226
8.6864 | -34.216 | -15.1327
8.6925 | -34.2163 | -15.1425
8.6986 | -34.2166 | -15.1526
8.7047 | -34.217 | -15.1625
8.7108 | -34.2173 | -15.1724
8.7169 | -34.2176 | -15.182
8.723 | -34.2179 | -15.1919
8.7291 | -34.2182 | -15.2016
8.7352 | -34.2186 | -15.2116
8.7413 | -34.2189 | -15.2214
8.7474 | -34.2192 | -15.2309
8.7535 | -34.2196 | -15.2408
8.7596 | -34.2199 | -15.2512
8.7657 | -34.2202 | -15.261
8.7718 | -34.2205 | -15.2705
8.7779 | -34.2209 | -15.2809
8.784 | -34.2211 | -15.2906
8.7901 | -34.2214 | -15.3008
8.7962 | -34.2218 | -15.3107
8.8023 | -34.2222 | -15.3202
8.8084 | -34.2225 | -15.3302
8.8146 | -34.2228 | -15.3399
8.8207 | -34.2231 | -15.3498
8.8268 | -34.2235 | -15.3595
8.8329 | -34.2238 | -15.3694
8.839 | -34.2242 | -15.3793
8.8451 | -34.2245 | -15.389
8.8512 | -34.2249 | -15.3989
8.8573 | -34.2252 | -15.4085
8.8634 | -34.2256 | -15.4182
8.8695 | -34.2259 | -15.4282
8.8756 | -34.2262 | -15.4378
8.8817 | -34.2265 | -15.4479
8.8878 | -34.2268 | -15.4579
8.8939 | -34.2271 | -15.4675
8.9 | -34.2274 | -15.477
8.9061 | -34.2278 | -15.4869
8.9122 | -34.2281 | -15.4966
8.9183 | -34.2285 | -15.5063
8.9244 | -34.2288 | -15.5159
8.9305 | -34.2292 | -15.5257
8.9366 | -34.2295 | -15.5354
8.9427 | -34.2299 | -15.5453
8.9488 | -34.2302 | -15.555
8.955 | -34.2305 | -15.5648
8.9611 | -34.2309 | -15.5747
8.9672 | -34.2313 | -15.5849
8.9733 | -34.2316 | -15.5946
8.9794 | -34.2319 | -15.6045
8.9855 | -34.2323 | -15.6142
8.9916 | -34.2326 | -15.6239
8.9977 | -34.233 | -15.6336
9.0038 | -34.2334 | -15.6436
9.0099 | -34.2337 | -15.6535
9.016 | -34.2341 | -15.6633
9.0221 | -34.2344 | -15.6733
9.0282 | -34.2348 | -15.6832
9.0343 | -34.2352 | -15.6933
9.0404 | -34.2356 | -15.7029
9.0465 | -34.2359 | -15.7131
9.0526 | -34.2363 | -15.7229
9.0587 | -34.2367 | -15.7326
9.0648 | -34.2371 | -15.7423
9.0709 | -34.2374 | -15.7519
9.077 | -34.2377 | -15.7622
9.0831 | -34.2381 | -15.7721
9.0892 | -34.2384 | -15.7819
9.0953 | -34.2387 | -15.7918
9.1015 | -34.2391 | -15.8016
9.1076 | -34.2395 | -15.8115
9.1137 | -34.2398 | -15.8212
9.1198 | -34.2402 | -15.8306
9.1259 | -34.2406 | -15.8405
9.132 | -34.241 | -15.8506
9.1381 | -34.2414 | -15.8603
9.1442 | -34.2418 | -15.8696
9.1503 | -34.2421 | -15.8796
9.1564 | -34.2425 | -15.8894
9.1625 | -34.2428 | -15.8993
9.1686 | -34.2432 | -15.9091
9.1747 | -34.2436 | -15.919
9.1808 | -34.244 | -15.929
9.1869 | -34.2443 | -15.9391
9.193 | -34.2447 | -15.9491
9.1991 | -34.245 | -15.9586
9.2052 | -34.2454 | -15.9683
9.2113 | -34.2458 | -15.9784
9.2174 | -34.2462 | -15.9888
9.2235 | -34.2466 | -15.9983
9.2296 | -34.247 | -16.0082
9.2357 | -34.2473 | -16.0181
9.2419 | -34.2477 | -16.0279
9.248 | -34.2481 | -16.0379
9.2541 | -34.2485 | -16.0479
9.2602 | -34.2488 | -16.0578
9.2663 | -34.2492 | -16.0675
9.2724 | -34.2495 | -16.0774
9.2785 | -34.2499 | -16.0871
9.2846 | -34.2503 | -16.097
9.2907 | -34.2507 | -16.1068
9.2968 | -34.251 | -16.1167
9.3029 | -34.2514 | -16.1267
9.309 | -34.2518 | -16.1364
9.3151 | -34.2521 | -16.1461
9.3212 | -34.2526 | -16.1561
9.3273 | -34.253 | -16.1658
9.3334 | -34.2534 | -16.1754
9.3395 | -34.2538 | -16.1851
9.3456 | -34.2542 | -16.1947
9.3517 | -34.2546 | -16.2041
9.3578 | -34.255 | -16.2138
9.3639 | -34.2554 | -16.2237
9.37 | -34.2558 | -16.2336
9.3761 | -34.2562 | -16.2433
9.3822 | -34.2566 | -16.2531
9.3884 | -34.257 | -16.2628
9.3945 | -34.2574 | -16.2727
9.4006 | -34.2578 | -16.2825
9.4067 | -34.2581 | -16.2925
9.4128 | -34.2585 | -16.3022
9.4189 | -34.259 | -16.3122
9.425 | -34.2594 | -16.3219
9.4311 | -34.2597 | -16.3317
9.4372 | -34.2601 | -16.3415
9.4433 | -34.2605 | -16.3514
9.4494 | -34.2609 | -16.3611
9.4555 | -34.2613 | -16.3709
9.4616 | -34.2617 | -16.3806
9.4677 | -34.2621 | -16.3902
9.4738 | -34.2625 | -16.4003
9.4799 | -34.2629 | -16.4101
9.486 | -34.2632 | -16.4202
9.4921 | -34.2636 | -16.4304
9.4982 | -34.264 | -16.4397
9.5043 | -34.2644 | -16.4492
9.5104 | -34.2648 | -16.4589
9.5165 | -34.2652 | -16.4687
9.5226 | -34.2656 | -16.4787
9.5288 | -34.266 | -16.4884
9.5349 | -34.2664 | -16.4982
9.541 | -34.2668 | -16.5079
9.5471 | -34.2672 | -16.5178
9.5532 | -34.2676 | -16.5275
9.5593 | -34.2681 | -16.5373
9.5654 | -34.2685 | -16.5471
9.5715 | -34.2689 | -16.5571
9.5776 | -34.2693 | -16.5672
9.5837 | -34.2696 | -16.5767
9.5898 | -34.27 | -16.5867
9.5959 | -34.2704 | -16.5964
9.602 | -34.2708 | -16.6061
9.6081 | -34.2711 | -16.616
9.6142 | -34.2716 | -16.6257
9.6203 | -34.2719 | -16.6355
9.6264 | -34.2723 | -16.6449
9.6325 | -34.2726 | -16.6549
9.6386 | -34.2731 | -16.6645
9.6447 | -34.2734 | -16.6742
9.6508 | -34.2739 | -16.6841
9.6569 | -34.2743 | -16.6937
9.663 | -34.2747 | -16.7036
9.6691 | -34.2751 | -16.7133
9.6753 | -34.2756 | -16.7231
9.6814 | -34.276 | -16.7327
9.6875 | -34.2764 | -16.7423
9.6936 | -34.2768 | -16.7519
9.6997 | -34.2772 | -16.7619
9.7058 | -34.2777 | -16.7717
9.7119 | -34.278 | -16.7814
9.718 | -34.2784 | -16.7914
9.7241 | -34.2788 | -16.8012
9.7302 | -34.2791 | -16.8111
9.7363 | -34.2796 | -16.8206
9.7424 | -34.28 | -16.8302
9.7485 | -34.2804 | -16.8398
9.7546 | -34.2808 | -16.8493
9.7607 | -34.2812 | -16.8592
9.7668 | -34.2816 | -16.8694
9.7729 | -34.2821 | -16.8788
9.779 | -34.2826 | -16.8885
9.7851 | -34.2829 | -16.8983
9.7912 | -34.2833 | -16.9086
9.7973 | -34.2837 | -16.9186
9.8034 | -34.2841 | -16.9285
9.8095 | -34.2845 | -16.9382
9.8157 | -34.2849 | -16.9479
9.8218 | -34.2853 | -16.9573
9.8279 | -34.2857 | -16.9672
9.834 | -34.2861 | -16.9768
9.8401 | -34.2866 | -16.9865
9.8462 | -34.287 | -16.9959
9.8523 | -34.2874 | -17.0056
9.8584 | -34.2879 | -17.0156
9.8645 | -34.2882 | -17.0254
9.8706 | -34.2886 | -17.035
9.8767 | -34.289 | -17.0444
9.8828 | -34.2895 | -17.0537
9.8889 | -34.2899 | -17.0637
9.895 | -34.2903 | -17.0731
9.9011 | -34.2906 | -17.0824
9.9072 | -34.2911 | -17.0917
9.9133 | -34.2915 | -17.1012
9.9194 | -34.2919 | -17.1107
9.9255 | -34.2923 | -17.1205
9.9316 | -34.2928 | -17.13
9.9377 | -34.2932 | -17.1395
9.9438 | -34.2936 | -17.1491
9.9499 | -34.294 | -17.1589
9.956 | -34.2944 | -17.1684
9.9622 | -34.2948 | -17.1779
9.9683 | -34.2952 | -17.1877
9.9744 | -34.2957 | -17.1973
9.9805 | -34.2961 | -17.207
9.9866 | -34.2965 | -17.2163
9.9927 | -34.2969 | -17.2262
9.9988 | -34.2974 | -17.2357
10.0049 | -34.2977 | -17.2453
10.011 | -34.2982 | -17.2549
10.0171 | -34.2986 | -17.2645
10.0232 | -34.299 | -17.2738
10.0293 | -34.2995 | -17.2833
10.0354 | -34.3 | -17.2926
10.0415 | -34.3004 | -17.3022
10.0476 | -34.3008 | -17.3122
10.0537 | -34.3011 | -17.3221
10.0598 | -34.3016 | -17.3318
10.0659 | -34.3019 | -17.3413
10.072 | -34.3024 | -17.3508
10.0781 | -34.3029 | -17.3604
10.0842 | -34.3033 | -17.3702
10.0903 | -34.3037 | -17.38
10.0964 | -34.3042 | -17.3894
10.1026 | -34.3046 | -17.399
10.1087 | -34.3051 | -17.4087
10.1148 | -34.3055 | -17.4184
10.1209 | -34.3059 | -17.4277
10.127 | -34.3063 | -17.4371
10.1331 | -34.3068 | -17.4465
10.1392 | -34.3072 | -17.456
10.1453 | -34.3076 | -17.4656
10.1514 | -34.308 | -17.4749
10.1575 | -34.3084 | -17.4844
10.1636 | -34.3089 | -17.4933
10.1697 | -34.3093 | -17.5029
10.1758 | -34.3097 | -17.5122
10.1819 | -34.3102 | -17.5219
10.188 | -34.3106 | -17.5312
10.1941 | -34.3109 | -17.5409
10.2002 | -34.3113 | -17.5498
10.2063 | -34.3117 | -17.5596
10.2124 | -34.3121 | -17.5691
10.2185 | -34.3126 | -17.5786
10.2246 | -34.313 | -17.5881
10.2307 | -34.3134 | -17.5978
10.2368 | -34.3138 | -17.6066
10.2429 | -34.3143 | -17.6158
10.2491 | -34.3147 | -17.625
10.2552 | -34.3151 | -17.6349
10.2613 | -34.3156 | -17.6447
10.2674 | -34.3159 | -17.6544
10.2735 | -34.3163 | -17.6641
10.2796 | -34.3167 | -17.6736
10.2857 | -34.3171 | -17.6832
10.2918 | -34.3175 | -17.6928
10.2979 | -34.318 | -17.7025
10.304 | -34.3185 | -17.7122
10.3101 | -34.3189 | -17.7216
10.3162 | -34.3193 | -17.7311
10.3223 | -34.3197 | -17.7406
10.3284 | -34.3202 | -17.75
10.3345 | -34.3206 | -17.7594
10.3406 | -34.321 | -17.7688
10.3467 | -34.3214 | -17.7782
10.3528 | -34.3218 | -17.7874
10.3589 | -34.3222 | -17.7971
10.365 | -34.3226 | -17.8069
10.3711 | -34.323 | -17.8166
10.3772 | -34.3233 | -17.8261
10.3833 | -34.3237 | -17.8355
10.3895 | -34.3242 | -17.8449
10.3956 | -34.3246 | -17.8542
10.4017 | -34.325 | -17.8635
10.4078 | -34.3254 | -17.8731
10.4139 | -34.3258 | -17.8825
10.42 | -34.3263 | -17.892
10.4261 | -34.3267 | -17.9014
10.4322 | -34.3272 | -17.9107
10.4383 | -34.3276 | -17.9202
10.4444 | -34.328 | -17.9294
10.4505 | -34.3284 | -17.9388
10.4566 | -34.3288 | -17.9483
10.4627 | -34.3292 | -17.9576
10.4688 | -34.3297 | -17.9666
10.4749 | -34.3301 | -17.9759
10.481 | -34.3306 | -17.9854
10.4871 | -34.331 | -17.9949
10.4932 | -34.3314 | -18.0043
10.4993 | -34.3319 | -18.0139
10.5054 | -34.3323 | -18.0233
10.5115 | -34.3327 | -18.0327
10.5176 | -34.3332 | -18.0418
10.5237 | -34.3336 | -18.0512
10.5298 | -34.334 | -18.0604
10.536 | -34.3345 | -18.07
10.5421 | -34.3349 | -18.079
10.5482 | -34.3353 | -18.0887
10.5543 | -34.3357 | -18.0981
10.5604 | -34.3362 | -18.1071
10.5665 | -34.3366 | -18.1163
10.5726 | -34.337 | -18.1254
10.5787 | -34.3374 | -18.135
10.5848 | -34.3379 | -18.144
10.5909 | -34.3383 | -18.1535
10.597 | -34.3387 | -18.1628
10.6031 | -34.3391 | -18.1723
10.6092 | -34.3395 | -18.1819
10.6153 | -34.3399 | -18.1915
10.6214 | -34.3404 | -18.2007
10.6275 | -34.3408 | -18.21
10.6336 | -34.3413 | -18.2197
10.6397 | -34.3417 | -18.2289
10.6458 | -34.342 | -18.238
10.6519 | -34.3424 | -18.2473
10.658 | -34.3428 | -18.2565
10.6641 | -34.3433 | -18.2659
10.6702 | -34.3437 | -18.2754
10.6764 | -34.3442 | -18.2846
10.6825 | -34.3446 | -18.2941
10.6886 | -34.3451 | -18.3034
10.6947 | -34.3455 | -18.3126
10.7008 | -34.3459 | -18.322
10.7069 | -34.3463 | -18.3314
10.713 | -34.3468 | -18.3407
10.7191 | -34.3472 | -18.3503
10.7252 | -34.3477 | -18.3595
10.7313 | -34.3481 | -18.3688
10.7374 | -34.3485 | -18.3779
10.7435 | -34.3489 | -18.3875
10.7496 | -34.3494 | -18.3967
10.7557 | -34.3498 | -18.406
10.7618 | -34.3502 | -18.4154
10.7679 | -34.3506 | -18.4246
10.774 | -34.351 | -18.434
10.7801 | -34.3514 | -18.4436
10.7862 | -34.3519 | -18.4528
10.7923 | -34.3523 | -18.4622
10.7984 | -34.3528 | -18.4714
10.8045 | -34.3532 | -18.4808
10.8106 | -34.3536 | -18.4901
10.8168 | -34.354 | -18.4995
10.8229 | -34.3544 | -18.5087
10.829 | -34.3549 | -18.5181
10.8351 | -34.3553 | -18.5275
10.8412 | -34.3557 | -18.5367
10.8473 | -34.3561 | -18.5456
10.8534 | -34.3566 | -18.5549
10.8595 | -34.357 | -18.5643
10.8656 | -34.3575 | -18.5733
10.8717 | -34.3579 | -18.5821
10.8778 | -34.3584 | -18.5913
10.8839 | -34.3588 | -18.6003
10.89 | -34.3592 | -18.6094
10.8961 | -34.3596 | -18.6187
10.9022 | -34.3601 | -18.628
10.9083 | -34.3606 | -18.6375
10.9144 | -34.361 | -18.6466
10.9205 | -34.3615 | -18.6557
10.9266 | -34.362 | -18.6652
10.9327 | -34.3625 | -18.6744
10.9388 | -34.3629 | -18.6835
10.9449 | -34.3633 | -18.6929
10.951 | -34.3637 | -18.702
10.9571 | -34.3641 | -18.7113
10.9633 | -34.3646 | -18.7208
10.9694 | -34.365 | -18.7304
10.9755 | -34.3655 | -18.7397
10.9816 | -34.3659 | -18.7491
10.9877 | -34.3663 | -18.7585
10.9938 | -34.3667 | -18.7681
10.9999 | -34.3672 | -18.7772
11.006 | -34.3677 | -18.7863
11.0121 | -34.3681 | -18.7959
11.0182 | -34.3685 | -18.8049
11.0243 | -34.3689 | -18.8142
11.0304 | -34.3694 | -18.8233
11.0365 | -34.3698 | -18.8325
11.0426 | -34.3702 | -18.8416
11.0487 | -34.3706 | -18.8508
11.0548 | -34.3711 | -18.8599
11.0609 | -34.3715 | -18.8694
11.067 | -34.372 | -18.8788
11.0731 | -34.3724 | -18.8881
11.0792 | -34.3729 | -18.8973
11.0853 | -34.3733 | -18.9063
11.0914 | -34.3737 | -18.9153
11.0975 | -34.3742 | -18.9248
11.1037 | -34.3746 | -18.9338
11.1098 | -34.375 | -18.9433
11.1159 | -34.3755 | -18.9527
11.122 | -34.3759 | -18.9622
11.1281 | -34.3763 | -18.9711
11.1342 | -34.3767 | -18.98
11.1403 | -34.3772 | -18.9896
11.1464 | -34.3776 | -18.9988
11.1525 | -34.3781 | -19.008
11.1586 | -34.3785 | -19.0177
11.1647 | -34.3789 | -19.0268
11.1708 | -34.3793 | -19.0358
11.1769 | -34.3797 | -19.0446
11.183 | -34.3802 | -19.0537
11.1891 | -34.3807 | -19.0634
11.1952 | -34.3811 | -19.073
11.2013 | -34.3815 | -19.082
11.2074 | -34.3819 | -19.0909
11.2135 | -34.3823 | -19.1002
11.2196 | -34.3827 | -19.1093
11.2257 | -34.3831 | -19.1184
11.2318 | -34.3835 | -19.1278
11.2379 | -34.3839 | -19.1369
11.244 | -34.3843 | -19.146
11.2502 | -34.3847 | -19.1552
11.2563 | -34.3852 | -19.1644
11.2624 | -34.3856 | -19.1735
11.2685 | -34.3861 | -19.1826
11.2746 | -34.3865 | -19.1914
11.2807 | -34.3869 | -19.2011
11.2868 | -34.3873 | -19.2107
11.2929 | -34.3878 | -19.2197
11.299 | -34.3882 | -19.2292
11.3051 | -34.3886 | -19.2384
11.3112 | -34.389 | -19.2476
11.3173 | -34.3894 | -19.2568
11.3234 | -34.3898 | -19.2659
11.3295 | -34.3902 | -19.2748
11.3356 | -34.3907 | -19.2839
11.3417 | -34.3911 | -19.2931
11.3478 | -34.3915 | -19.3024
11.3539 | -34.3919 | -19.3114
11.36 | -34.3923 | -19.3205
11.3661 | -34.3928 | -19.3297
11.3722 | -34.3932 | -19.3389
11.3783 | -34.3936 | -19.348
11.3844 | -34.394 | -19.3572
11.3906 | -34.3944 | -19.3667
11.3967 | -34.3948 | -19.3757
11.4028 | -34.3952 | -19.385
11.4089 | -34.3957 | -19.3941
11.415 | -34.3961 | -19.4032
11.4211 | -34.3965 | -19.4123
11.4272 | -34.3969 | -19.4215
11.4333 | -34.3973 | -19.4306
11.4394 | -34.3978 | -19.4402
11.4455 | -34.3981 | -19.4493
11.4516 | -34.3986 | -19.4581
11.4577 | -34.399 | -19.4675
11.4638 | -34.3995 | -19.4763
11.4699 | -34.3999 | -19.4859
11.476 | -34.4003 | -19.4955
11.4821 | -34.4007 | -19.5043
11.4882 | -34.4012 | -19.5138
11.4943 | -34.4016 | -19.5228
11.5004 | -34.4021 | -19.5319
11.5065 | -34.4025 | -19.5411
11.5126 | -34.403 | -19.5498
11.5187 | -34.4035 | -19.5591
11.5248 | -34.4039 | -19.5683
11.5309 | -34.4043 | -19.5774
11.5371 | -34.4047 | -19.5864
11.5432 | -34.4051 | -19.5957
11.5493 | -34.4055 | -19.605
11.5554 | -34.4059 | -19.614
11.5615 | -34.4064 | -19.6233
11.5676 | -34.4068 | -19.6327
11.5737 | -34.4072 | -19.6415
11.5798 | -34.4077 | -19.6504
11.5859 | -34.4081 | -19.6594
11.592 | -34.4085 | -19.6684
11.5981 | -34.409 | -19.6774
11.6042 | -34.4094 | -19.6866
11.6103 | -34.4098 | -19.6956
11.6164 | -34.4103 | -19.7044
11.6225 | -34.4107 | -19.7134
11.6286 | -34.4111 | -19.7226
11.6347 | -34.4115 | -19.7318
11.6408 | -34.412 | -19.7409
11.6469 | -34.4125 | -19.7503
11.653 | -34.4129 | -19.7592
11.6591 | -34.4133 | -19.7683
11.6652 | -34.4137 | -19.7775
11.6713 | -34.4141 | -19.7866
11.6775 | -34.4146 | -19.7956
11.6836 | -34.415 | -19.8048
11.6897 | -34.4155 | -19.8138
11.6958 | -34.4159 | -19.8233
11.7019 | -34.4164 | -19.8324
11.708 | -34.4168 | -19.8418
11.7141 | -34.4173 | -19.8509
11.7202 | -34.4177 | -19.8597
11.7263 | -34.4181 | -19.8689
11.7324 | -34.4185 | -19.8781
11.7385 | -34.4189 | -19.8874
11.7446 | -34.4194 | -19.8967
11.7507 | -34.4198 | -19.9058
11.7568 | -34.4202 | -19.9149
11.7629 | -34.4207 | -19.9241
11.769 | -34.4211 | -19.9333
11.7751 | -34.4216 | -19.9421
11.7812 | -34.422 | -19.9511
11.7873 | -34.4224 | -19.9605
11.7934 | -34.4229 | -19.9696
11.7995 | -34.4233 | -19.9783
11.8056 | -34.4237 | -19.9878
11.8117 | -34.4241 | -19.997
11.8178 | -34.4245 | -20.006
11.824 | -34.425 | -20.0148
11.8301 | -34.4255 | -20.0236
11.8362 | -34.4259 | -20.0322
11.8423 | -34.4263 | -20.0413
11.8484 | -34.4267 | -20.0502
11.8545 | -34.4271 | -20.0594
11.8606 | -34.4275 | -20.0687
11.8667 | -34.428 | -20.0777
11.8728 | -34.4284 | -20.0867
11.8789 | -34.4288 | -20.0961
11.885 | -34.4292 | -20.1049
11.8911 | -34.4297 | -20.1142
11.8972 | -34.4301 | -20.1231
11.9033 | -34.4305 | -20.1317
11.9094 | -34.4309 | -20.1404
11.9155 | -34.4313 | -20.1494
11.9216 | -34.4318 | -20.1583
11.9277 | -34.4323 | -20.1671
11.9338 | -34.4326 | -20.1762
11.9399 | -34.4331 | -20.1855
11.946 | -34.4335 | -20.1945
11.9521 | -34.4339 | -20.2036
11.9582 | -34.4344 | -20.2125
11.9644 | -34.4348 | -20.2214
11.9705 | -34.4353 | -20.2305
11.9766 | -34.4357 | -20.2395
11.9827 | -34.4361 | -20.2484
11.9888 | -34.4366 | -20.2573
11.9949 | -34.4369 | -20.2662
12.001 | -34.4374 | -20.2753
12.0071 | -34.4379 | -20.284
12.0132 | -34.4383 | -20.2927
12.0193 | -34.4387 | -20.302
12.0254 | -34.4391 | -20.3107
12.0315 | -34.4396 | -20.3195
12.0376 | -34.44 | -20.3283
12.0437 | -34.4404 | -20.3371
12.0498 | -34.4409 | -20.3463
12.0559 | -34.4413 | -20.3553
12.062 | -34.4418 | -20.3641
12.0681 | -34.4422 | -20.3731
12.0742 | -34.4426 | -20.3822
12.0803 | -34.4431 | -20.3912
12.0864 | -34.4435 | -20.4
12.0925 | -34.444 | -20.409
12.0986 | -34.4444 | -20.4178
12.1047 | -34.4449 | -20.4266
12.1109 | -34.4453 | -20.4358
12.117 | -34.4458 | -20.4447
12.1231 | -34.4462 | -20.4536
12.1292 | -34.4466 | -20.4628
12.1353 | -34.4471 | -20.4714
12.1414 | -34.4476 | -20.4805
12.1475 | -34.448 | -20.4896
12.1536 | -34.4484 | -20.4988
12.1597 | -34.4488 | -20.5077
12.1658 | -34.4493 | -20.5166
12.1719 | -34.4498 | -20.5258
12.178 | -34.4502 | -20.535
12.1841 | -34.4507 | -20.5441
12.1902 | -34.4511 | -20.5531
12.1963 | -34.4516 | -20.562
12.2024 | -34.452 | -20.5709
12.2085 | -34.4524 | -20.5799
12.2146 | -34.4529 | -20.5886
12.2207 | -34.4533 | -20.5975
12.2268 | -34.4538 | -20.6064
12.2329 | -34.4542 | -20.6154
12.239 | -34.4546 | -20.6239
12.2451 | -34.4551 | -20.6326
12.2513 | -34.4556 | -20.6416
12.2574 | -34.456 | -20.6507
12.2635 | -34.4565 | -20.6593
12.2696 | -34.4569 | -20.6681
12.2757 | -34.4573 | -20.6767
12.2818 | -34.4577 | -20.6854
12.2879 | -34.4581 | -20.6941
12.294 | -34.4585 | -20.7029
12.3001 | -34.459 | -20.7119
12.3062 | -34.4594 | -20.7207
12.3123 | -34.4599 | -20.7297
12.3184 | -34.4603 | -20.7384
12.3245 | -34.4608 | -20.7473
12.3306 | -34.4612 | -20.7558
12.3367 | -34.4617 | -20.7645
12.3428 | -34.4621 | -20.7731
12.3489 | -34.4626 | -20.7813
12.355 | -34.463 | -20.7902
12.3611 | -34.4634 | -20.7989
12.3672 | -34.4639 | -20.8076
12.3733 | -34.4643 | -20.8168
12.3794 | -34.4647 | -20.8251
12.3855 | -34.4651 | -20.834
12.3916 | -34.4656 | -20.8428
12.3978 | -34.466 | -20.8519
12.4039 | -34.4664 | -20.8607
12.41 | -34.4669 | -20.8697
12.4161 | -34.4673 | -20.878
12.4222 | -34.4678 | -20.8867
12.4283 | -34.4683 | -20.8956
12.4344 | -34.4687 | -20.9043
12.4405 | -34.4691 | -20.9132
12.4466 | -34.4695 | -20.9218
12.4527 | -34.47 | -20.9306
12.4588 | -34.4704 | -20.9394
12.4649 | -34.4709 | -20.9485
12.471 | -34.4713 | -20.9577
12.4771 | -34.4718 | -20.9659
12.4832 | -34.4722 | -20.9746
12.4893 | -34.4727 | -20.9831
12.4954 | -34.4732 | -20.9917
12.5015 | -34.4737 | -21.0002
12.5076 | -34.474 | -21.0091
12.5137 | -34.4744 | -21.0182
12.5198 | -34.4749 | -21.0265
12.5259 | -34.4754 | -21.0352
12.532 | -34.4758 | -21.0442
12.5382 | -34.4763 | -21.0529
12.5443 | -34.4767 | -21.0616
12.5504 | -34.4771 | -21.0704
12.5565 | -34.4776 | -21.0788
12.5626 | -34.4781 | -21.0877
12.5687 | -34.4785 | -21.0965
12.5748 | -34.479 | -21.1052
12.5809 | -34.4794 | -21.1136
12.587 | -34.4798 | -21.1228
12.5931 | -34.4803 | -21.1314
12.5992 | -34.4807 | -21.1403
12.6053 | -34.4812 | -21.149
12.6114 | -34.4817 | -21.1578
12.6175 | -34.4821 | -21.1666
12.6236 | -34.4825 | -21.175
12.6297 | -34.483 | -21.1837
12.6358 | -34.4835 | -21.1921
12.6419 | -34.484 | -21.2009
12.648 | -34.4844 | -21.21
12.6541 | -34.4848 | -21.2184
12.6602 | -34.4853 | -21.2267
12.6663 | -34.4856 | -21.2355
12.6724 | -34.4861 | -21.2442
12.6785 | -34.4865 | -21.2532
12.6847 | -34.487 | -21.2618
12.6908 | -34.4875 | -21.2706
12.6969 | -34.488 | -21.2795
12.703 | -34.4885 | -21.2882
12.7091 | -34.4889 | -21.2971
12.7152 | -34.4893 | -21.3059
12.7213 | -34.4897 | -21.3147
12.7274 | -34.4902 | -21.3234
12.7335 | -34.4907 | -21.3314
12.7396 | -34.4911 | -21.3396
12.7457 | -34.4916 | -21.3482
12.7518 | -34.492 | -21.357
12.7579 | -34.4925 | -21.3658
12.764 | -34.4929 | -21.3744
12.7701 | -34.4932 | -21.3827
12.7762 | -34.4937 | -21.3913
12.7823 | -34.4941 | -21.4001
12.7884 | -34.4945 | -21.4087
12.7945 | -34.4949 | -21.417
12.8006 | -34.4954 | -21.4248
12.8067 | -34.4959 | -21.4338
12.8128 | -34.4963 | -21.442
12.8189 | -34.4967 | -21.4503
12.8251 | -34.4971 | -21.4587
12.8312 | -34.4975 | -21.4674
12.8373 | -34.498 | -21.476
12.8434 | -34.4985 | -21.4845
12.8495 | -34.4989 | -21.4931
12.8556 | -34.4993 | -21.5015
12.8617 | -34.4997 | -21.5099
12.8678 | -34.5001 | -21.5185
12.8739 | -34.5006 | -21.5268
12.88 | -34.501 | -21.5355
12.8861 | -34.5014 | -21.5443
12.8922 | -34.5018 | -21.5529
12.8983 | -34.5023 | -21.561
12.9044 | -34.5028 | -21.5694
12.9105 | -34.5033 | -21.5779
12.9166 | -34.5037 | -21.5869
12.9227 | -34.5042 | -21.5959
12.9288 | -34.5047 | -21.6046
12.9349 | -34.5051 | -21.6134
12.941 | -34.5056 | -21.6221
12.9471 | -34.506 | -21.6306
12.9532 | -34.5065 | -21.6391
12.9593 | -34.507 | -21.6478
12.9654 | -34.5074 | -21.6566
12.9716 | -34.5078 | -21.6654
12.9777 | -34.5083 | -21.674
12.9838 | -34.5087 | -21.6825
12.9899 | -34.5091 | -21.6911
12.996 | -34.5096 | -21.7
13.0021 | -34.51 | -21.7088
13.0082 | -34.5104 | -21.7174
13.0143 | -34.5108 | -21.7251
13.0204 | -34.5112 | -21.7332
13.0265 | -34.5117 | -21.7414
13.0326 | -34.5121 | -21.7502
13.0387 | -34.5126 | -21.7587
13.0448 | -34.513 | -21.7671
13.0509 | -34.5134 | -21.7757
13.057 | -34.5139 | -21.7837
13.0631 | -34.5144 | -21.7918
13.0692 | -34.5149 | -21.8001
13.0753 | -34.5154 | -21.8085
13.0814 | -34.5158 | -21.8167
13.0875 | -34.5162 | -21.8248
13.0936 | -34.5167 | -21.8332
13.0997 | -34.5171 | -21.8412
13.1058 | -34.5175 | -21.8491
13.112 | -34.5179 | -21.8576
13.1181 | -34.5184 | -21.866
13.1242 | -34.5188 | -21.8745
13.1303 | -34.5193 | -21.8831
13.1364 | -34.5197 | -21.8915
13.1425 | -34.5201 | -21.8999
13.1486 | -34.5206 | -21.9081
13.1547 | -34.5211 | -21.9164
13.1608 | -34.5215 | -21.9248
13.1669 | -34.5219 | -21.9333
13.173 | -34.5224 | -21.9417
13.1791 | -34.5228 | -21.9503
13.1852 | -34.5232 | -21.9589
13.1913 | -34.5237 | -21.9672
13.1974 | -34.5241 | -21.9758
13.2035 | -34.5245 | -21.9843
13.2096 | -34.525 | -21.9927
13.2157 | -34.5254 | -22.0011
13.2218 | -34.5259 | -22.0096
13.2279 | -34.5264 | -22.0179
13.234 | -34.5268 | -22.0264
13.2401 | -34.5273 | -22.0349
13.2462 | -34.5277 | -22.0432
13.2524 | -34.5282 | -22.0514
13.2585 | -34.5286 | -22.0597
13.2646 | -34.529 | -22.0683
13.2707 | -34.5295 | -22.0765
13.2768 | -34.5299 | -22.0854
13.2829 | -34.5303 | -22.0938
13.289 | -34.5308 | -22.1022
13.2951 | -34.5313 | -22.1103
13.3012 | -34.5317 | -22.1188
13.3073 | -34.5321 | -22.1274
13.3134 | -34.5325 | -22.1358
13.3195 | -34.533 | -22.144
13.3256 | -34.5335 | -22.1524
13.3317 | -34.5339 | -22.1606
13.3378 | -34.5343 | -22.1691
13.3439 | -34.5348 | -22.1774
13.35 | -34.5352 | -22.1858
13.3561 | -34.5357 | -22.1943
13.3622 | -34.5361 | -22.2026
13.3683 | -34.5366 | -22.2108
13.3744 | -34.537 | -22.2191
13.3805 | -34.5374 | -22.2275
13.3866 | -34.5378 | -22.2357
13.3927 | -34.5383 | -22.244
13.3989 | -34.5387 | -22.2521
13.405 | -34.5392 | -22.2603
13.4111 | -34.5397 | -22.2686
13.4172 | -34.5401 | -22.2767
13.4233 | -34.5405 | -22.2854
13.4294 | -34.541 | -22.2935
13.4355 | -34.5414 | -22.3026
13.4416 | -34.5418 | -22.3111
13.4477 | -34.5422 | -22.3195
13.4538 | -34.5427 | -22.3279
13.4599 | -34.5431 | -22.3365
13.466 | -34.5436 | -22.3451
13.4721 | -34.5441 | -22.3537
13.4782 | -34.5445 | -22.362
13.4843 | -34.545 | -22.3703
13.4904 | -34.5454 | -22.3782
13.4965 | -34.5458 | -22.3869
13.5026 | -34.5462 | -22.3951
13.5087 | -34.5467 | -22.4031
13.5148 | -34.5471 | -22.4115
13.5209 | -34.5475 | -22.4197
13.527 | -34.548 | -22.4284
13.5331 | -34.5484 | -22.4367
13.5393 | -34.5488 | -22.4451
13.5454 | -34.5493 | -22.4535
13.5515 | -34.5497 | -22.462
13.5576 | -34.5502 | -22.4705
13.5637 | -34.5506 | -22.4787
13.5698 | -34.5511 | -22.4873
13.5759 | -34.5515 | -22.4954
13.582 | -34.5518 | -22.5037
13.5881 | -34.5523 | -22.5122
13.5942 | -34.5528 | -22.5206
13.6003 | -34.5532 | -22.5289
13.6064 | -34.5537 | -22.5375
13.6125 | -34.5541 | -22.5459
13.6186 | -34.5546 | -22.5541
13.6247 | -34.555 | -22.5625
13.6308 | -34.5554 | -22.5707
13.6369 | -34.5558 | -22.579
13.643 | -34.5563 | -22.5872
13.6491 | -34.5568 | -22.5957
13.6552 | -34.5572 | -22.604
13.6613 | -34.5577 | -22.6125
13.6674 | -34.5581 | -22.6209
13.6735 | -34.5585 | -22.6293
13.6796 | -34.5589 | -22.6374
13.6858 | -34.5593 | -22.6458
13.6919 | -34.5598 | -22.6543
13.698 | -34.5602 | -22.6626
13.7041 | -34.5606 | -22.6713
13.7102 | -34.5611 | -22.6796
13.7163 | -34.5616 | -22.6882
13.7224 | -34.562 | -22.6965
13.7285 | -34.5625 | -22.7051
13.7346 | -34.5629 | -22.7132
13.7407 | -34.5633 | -22.7216
13.7468 | -34.5637 | -22.7299
13.7529 | -34.5641 | -22.7382
13.759 | -34.5646 | -22.7468
13.7651 | -34.565 | -22.7553
13.7712 | -34.5654 | -22.7639
13.7773 | -34.5659 | -22.7723
13.7834 | -34.5664 | -22.7806
13.7895 | -34.5669 | -22.7887
13.7956 | -34.5673 | -22.7972
13.8017 | -34.5677 | -22.8053
13.8078 | -34.5682 | -22.8137
13.8139 | -34.5687 | -22.8218
13.82 | -34.5692 | -22.8299
13.8262 | -34.5696 | -22.8386
13.8323 | -34.5701 | -22.847
13.8384 | -34.5705 | -22.8553
13.8445 | -34.571 | -22.8637
13.8506 | -34.5715 | -22.8722
13.8567 | -34.5719 | -22.8808
13.8628 | -34.5723 | -22.8891
13.8689 | -34.5727 | -22.8975
13.875 | -34.5732 | -22.9058
13.8811 | -34.5736 | -22.9142
13.8872 | -34.5741 | -22.9226
13.8933 | -34.5745 | -22.9309
13.8994 | -34.575 | -22.9391
13.9055 | -34.5754 | -22.9475
13.9116 | -34.5759 | -22.9557
13.9177 | -34.5764 | -22.9637
13.9238 | -34.5769 | -22.9722
13.9299 | -34.5774 | -22.9806
13.936 | -34.5778 | -22.9889
13.9421 | -34.5782 | -22.9972
13.9482 | -34.5787 | -23.0056
13.9543 | -34.5791 | -23.0139
13.9604 | -34.5796 | -23.0224
13.9665 | -34.58 | -23.0309
13.9727 | -34.5805 | -23.039
13.9788 | -34.5809 | -23.0476
13.9849 | -34.5813 | -23.0562
13.991 | -34.5817 | -23.0647
13.9971 | -34.5821 | -23.0735
14.0032 | -34.5825 | -23.082
14.0093 | -34.5829 | -23.0903
14.0154 | -34.5834 | -23.0987
14.0215 | -34.5838 | -23.1068
14.0276 | -34.5843 | -23.115
14.0337 | -34.5848 | -23.1237
14.0398 | -34.5852 | -23.1318
14.0459 | -34.5856 | -23.1403
14.052 | -34.586 | -23.1487
14.0581 | -34.5864 | -23.1568
14.0642 | -34.5869 | -23.1653
14.0703 | -34.5873 | -23.1736
14.0764 | -34.5878 | -23.1824
14.0825 | -34.5883 | -23.1908
14.0886 | -34.5887 | -23.1987
14.0947 | -34.5891 | -23.2067
14.1008 | -34.5895 | -23.2154
14.1069 | -34.59 | -23.2237
14.1131 | -34.5904 | -23.2323
14.1192 | -34.5908 | -23.2404
14.1253 | -34.5912 | -23.2485
14.1314 | -34.5917 | -23.2568
14.1375 | -34.5922 | -23.2654
14.1436 | -34.5927 | -23.2738
14.1497 | -34.5932 | -23.2818
14.1558 | -34.5936 | -23.2903
14.1619 | -34.594 | -23.2987
14.168 | -34.5945 | -23.3071
14.1741 | -34.5949 | -23.3153
14.1802 | -34.5954 | -23.3236
14.1863 | -34.5958 | -23.332
14.1924 | -34.5963 | -23.3404
14.1985 | -34.5967 | -23.3485
14.2046 | -34.5972 | -23.3567
14.2107 | -34.5976 | -23.3648
14.2168 | -34.5981 | -23.3735
14.2229 | -34.5985 | -23.3818
14.229 | -34.599 | -23.3902
14.2351 | -34.5995 | -23.3982
14.2412 | -34.5999 | -23.4063
14.2473 | -34.6004 | -23.4148
14.2534 | -34.6009 | -23.4231
14.2596 | -34.6013 | -23.4314
14.2657 | -34.6017 | -23.4399
14.2718 | -34.6022 | -23.4484
14.2779 | -34.6026 | -23.4565
14.284 | -34.6031 | -23.4649
14.2901 | -34.6036 | -23.473
14.2962 | -34.6041 | -23.4811
14.3023 | -34.6045 | -23.4895
14.3084 | -34.6049 | -23.4982
14.3145 | -34.6054 | -23.5065
14.3206 | -34.6058 | -23.5149
14.3267 | -34.6063 | -23.5232
14.3328 | -34.6068 | -23.5314
14.3389 | -34.6072 | -23.5394
14.345 | -34.6076 | -23.5478
14.3511 | -34.6081 | -23.5559
14.3572 | -34.6085 | -23.5644
14.3633 | -34.6089 | -23.5726
14.3694 | -34.6093 | -23.5809
14.3755 | -34.6098 | -23.5889
14.3816 | -34.6102 | -23.5971
14.3877 | -34.6107 | -23.6054
14.3938 | -34.6111 | -23.614
14.4 | -34.6116 | -23.6221
14.4061 | -34.612 | -23.6304
14.4122 | -34.6124 | -23.639
14.4183 | -34.6128 | -23.6473
14.4244 | -34.6133 | -23.6557
14.4305 | -34.6137 | -23.6642
14.4366 | -34.6142 | -23.6727
14.4427 | -34.6147 | -23.6813
14.4488 | -34.6151 | -23.6893
14.4549 | -34.6155 | -23.6976
14.461 | -34.616 | -23.706
14.4671 | -34.6164 | -23.7144
14.4732 | -34.6169 | -23.7224
14.4793 | -34.6173 | -23.7314
14.4854 | -34.6179 | -23.7393
14.4915 | -34.6183 | -23.7474
14.4976 | -34.6188 | -23.7559
14.5037 | -34.6193 | -23.7643
14.5098 | -34.6197 | -23.7724
14.5159 | -34.6202 | -23.7809
14.522 | -34.6206 | -23.7894
14.5281 | -34.6211 | -23.7975
14.5342 | -34.6216 | -23.8059
14.5403 | -34.622 | -23.8142
14.5465 | -34.6225 | -23.8229
14.5526 | -34.623 | -23.8315
14.5587 | -34.6234 | -23.8399
14.5648 | -34.6238 | -23.848
14.5709 | -34.6243 | -23.8565
14.577 | -34.6247 | -23.8648
14.5831 | -34.6252 | -23.8733
14.5892 | -34.6257 | -23.8812
14.5953 | -34.6262 | -23.8892
14.6014 | -34.6266 | -23.8975
14.6075 | -34.6271 | -23.9056
14.6136 | -34.6275 | -23.9138
14.6197 | -34.6279 | -23.9222
14.6258 | -34.6284 | -23.9305
14.6319 | -34.6289 | -23.9386
14.638 | -34.6293 | -23.9469
14.6441 | -34.6298 | -23.9548
14.6502 | -34.6302 | -23.9628
14.6563 | -34.6307 | -23.9711
14.6624 | -34.6312 | -23.9793
14.6685 | -34.6316 | -23.9873
14.6746 | -34.6321 | -23.9957
14.6807 | -34.6325 | -24.004
14.6869 | -34.633 | -24.0122
14.693 | -34.6334 | -24.0203
14.6991 | -34.6338 | -24.0284
14.7052 | -34.6343 | -24.0366
14.7113 | -34.6348 | -24.045
14.7174 | -34.6352 | -24.0533
14.7235 | -34.6357 | -24.0614
14.7296 | -34.6361 | -24.0696
14.7357 | -34.6366 | -24.0777
14.7418 | -34.6371 | -24.0858
14.7479 | -34.6375 | -24.0941
14.754 | -34.638 | -24.1023
14.7601 | -34.6385 | -24.1107
14.7662 | -34.6389 | -24.119
14.7723 | -34.6394 | -24.1271
14.7784 | -34.6398 | -24.1354
14.7845 | -34.6403 | -24.1438
14.7906 | -34.6408 | -24.1521
14.7967 | -34.6412 | -24.1602
14.8028 | -34.6416 | -24.1683
14.8089 | -34.6421 | -24.1765
14.815 | -34.6425 | -24.1845
14.8211 | -34.643 | -24.1926
14.8272 | -34.6434 | -24.201
14.8334 | -34.6439 | -24.2091
14.8395 | -34.6443 | -24.2173
14.8456 | -34.6448 | -24.2254
14.8517 | -34.6453 | -24.2337
14.8578 | -34.6457 | -24.2418
14.8639 | -34.6462 | -24.2502
14.87 | -34.6466 | -24.2586
14.8761 | -34.6471 | -24.2665
14.8822 | -34.6475 | -24.2748
14.8883 | -34.648 | -24.2829
14.8944 | -34.6485 | -24.2911
14.9005 | -34.6489 | -24.2992
14.9066 | -34.6494 | -24.3074
14.9127 | -34.6498 | -24.3155
14.9188 | -34.6503 | -24.3239
14.9249 | -34.6507 | -24.332
14.931 | -34.6512 | -24.34
14.9371 | -34.6517 | -24.3482
14.9432 | -34.6521 | -24.3565
14.9493 | -34.6526 | -24.3647
14.9554 | -34.653 | -24.3729
14.9615 | -34.6535 | -24.3811
14.9676 | -34.654 | -24.3895
14.9738 | -34.6544 | -24.3975
14.9799 | -34.6549 | -24.4061
14.986 | -34.6554 | -24.4138
14.9921 | -34.6559 | -24.4218
14.9982 | -34.6563 | -24.4304
15.0043 | -34.6567 | -24.4386
15.0104 | -34.6572 | -24.4467
15.0165 | -34.6577 | -24.4551
15.0226 | -34.6581 | -24.4629
15.0287 | -34.6586 | -24.4716
15.0348 | -34.6591 | -24.4799
15.0409 | -34.6595 | -24.4883
15.047 | -34.66 | -24.4968
15.0531 | -34.6606 | -24.5051
15.0592 | -34.661 | -24.5135
15.0653 | -34.6615 | -24.5217
15.0714 | -34.6619 | -24.5297
15.0775 | -34.6624 | -24.5375
15.0836 | -34.6629 | -24.5459
15.0897 | -34.6634 | -24.5541
15.0958 | -34.6639 | -24.5621
15.1019 | -34.6643 | -24.5703
15.108 | -34.6647 | -24.5785
15.1141 | -34.6652 | -24.5867
15.1203 | -34.6657 | -24.595
15.1264 | -34.6662 | -24.6034
15.1325 | -34.6667 | -24.6114
15.1386 | -34.6672 | -24.6198
15.1447 | -34.6677 | -24.6279
15.1508 | -34.6682 | -24.6363
15.1569 | -34.6687 | -24.6447
15.163 | -34.6692 | -24.6525
15.1691 | -34.6697 | -24.6607
15.1752 | -34.6702 | -24.6689
15.1813 | -34.6707 | -24.6771
15.1874 | -34.6712 | -24.6855
15.1935 | -34.6717 | -24.6936
15.1996 | -34.6722 | -24.702
15.2057 | -34.6727 | -24.7099
15.2118 | -34.6731 | -24.7179
15.2179 | -34.6736 | -24.7259
15.224 | -34.6741 | -24.7342
15.2301 | -34.6746 | -24.7421
15.2362 | -34.6751 | -24.7502
15.2423 | -34.6755 | -24.7585
15.2484 | -34.676 | -24.7665
15.2545 | -34.6765 | -24.7746
15.2607 | -34.677 | -24.7827
15.2668 | -34.6775 | -24.7907
15.2729 | -34.678 | -24.7988
15.279 | -34.6785 | -24.807
15.2851 | -34.679 | -24.8151
15.2912 | -34.6794 | -24.8232
15.2973 | -34.6799 | -24.8312
15.3034 | -34.6804 | -24.839
15.3095 | -34.6809 | -24.847
15.3156 | -34.6813 | -24.8552
15.3217 | -34.6818 | -24.8631
15.3278 | -34.6823 | -24.8714
15.3339 | -34.6828 | -24.8792
15.34 | -34.6833 | -24.8872
15.3461 | -34.6838 | -24.8954
15.3522 | -34.6842 | -24.9035
15.3583 | -34.6847 | -24.9116
15.3644 | -34.6852 | -24.9198
15.3705 | -34.6857 | -24.9276
15.3766 | -34.6862 | -24.9359
15.3827 | -34.6867 | -24.9439
15.3888 | -34.6871 | -24.9517
15.3949 | -34.6876 | -24.9596
15.401 | -34.6881 | -24.9677
15.4072 | -34.6886 | -24.9754
15.4133 | -34.6891 | -24.9835
15.4194 | -34.6895 | -24.9918
15.4255 | -34.69 | -25
15.4316 | -34.6905 | -25.0079
15.4377 | -34.691 | -25.016
15.4438 | -34.6915 | -25.0237
15.4499 | -34.6919 | -25.0317
15.456 | -34.6924 | -25.0398
15.4621 | -34.6929 | -25.0479
15.4682 | -34.6934 | -25.0562
15.4743 | -34.6939 | -25.0641
15.4804 | -34.6944 | -25.0723
15.4865 | -34.6949 | -25.0804
15.4926 | -34.6954 | -25.0885
15.4987 | -34.6959 | -25.0966
15.5048 | -34.6963 | -25.1046
15.5109 | -34.6968 | -25.1126
15.517 | -34.6973 | -25.1209
15.5231 | -34.6978 | -25.129
15.5292 | -34.6983 | -25.137
15.5353 | -34.6988 | -25.145
15.5414 | -34.6993 | -25.1534
15.5476 | -34.6998 | -25.1614
15.5537 | -34.7003 | -25.1694
15.5598 | -34.7008 | -25.1771
15.5659 | -34.7013 | -25.1854
15.572 | -34.7018 | -25.1931
15.5781 | -34.7023 | -25.2012
15.5842 | -34.7028 | -25.209
15.5903 | -34.7033 | -25.2167
15.5964 | -34.7038 | -25.2245
15.6025 | -34.7043 | -25.2323
15.6086 | -34.7048 | -25.2404
15.6147 | -34.7052 | -25.2483
15.6208 | -34.7058 | -25.2561
15.6269 | -34.7063 | -25.2639
15.633 | -34.7068 | -25.272
15.6391 | -34.7073 | -25.2802
15.6452 | -34.7078 | -25.2879
15.6513 | -34.7083 | -25.2961
15.6574 | -34.7088 | -25.3039
15.6635 | -34.7093 | -25.3119
15.6696 | -34.7098 | -25.3196
15.6757 | -34.7104 | -25.3281
15.6818 | -34.7108 | -25.3363
15.688 | -34.7114 | -25.3443
15.6941 | -34.7119 | -25.3523
15.7002 | -34.7123 | -25.3603
15.7063 | -34.7128 | -25.3685
15.7124 | -34.7133 | -25.3765
15.7185 | -34.7138 | -25.3843
15.7246 | -34.7143 | -25.3922
15.7307 | -34.7148 | -25.4004
15.7368 | -34.7153 | -25.4086
15.7429 | -34.7158 | -25.4164
15.749 | -34.7163 | -25.4245
15.7551 | -34.7168 | -25.4324
15.7612 | -34.7173 | -25.4398
15.7673 | -34.7177 | -25.4478
15.7734 | -34.7182 | -25.4556
15.7795 | -34.7187 | -25.464
15.7856 | -34.7192 | -25.4716
15.7917 | -34.7197 | -25.4794
15.7978 | -34.7203 | -25.4872
15.8039 | -34.7208 | -25.4951
15.81 | -34.7213 | -25.5029
15.8161 | -34.7217 | -25.5109
15.8222 | -34.7223 | -25.5188
15.8283 | -34.7228 | -25.5264
15.8345 | -34.7233 | -25.5343
15.8406 | -34.7238 | -25.5422
15.8467 | -34.7243 | -25.5502
15.8528 | -34.7248 | -25.5582
15.8589 | -34.7253 | -25.5661
15.865 | -34.7258 | -25.5737
15.8711 | -34.7263 | -25.5819
15.8772 | -34.7269 | -25.5899
15.8833 | -34.7274 | -25.5982
15.8894 | -34.7279 | -25.606
15.8955 | -34.7284 | -25.614
15.9016 | -34.7289 | -25.6219
15.9077 | -34.7294 | -25.6298
15.9138 | -34.7299 | -25.6376
15.9199 | -34.7304 | -25.6456
15.926 | -34.7309 | -25.6539
15.9321 | -34.7314 | -25.6619
15.9382 | -34.732 | -25.6698
15.9443 | -34.7325 | -25.6777
15.9504 | -34.733 | -25.6857
15.9565 | -34.7335 | -25.6937
15.9626 | -34.734 | -25.7016
15.9687 | -34.7345 | -25.7092
15.9749 | -34.7351 | -25.7172
15.981 | -34.7356 | -25.725
15.9871 | -34.7361 | -25.7329
15.9932 | -34.7366 | -25.741
15.9993 | -34.737 | -25.7488
16.0054 | -34.7376 | -25.7565
16.0115 | -34.7381 | -25.7642
16.0176 | -34.7386 | -25.7723
16.0237 | -34.7391 | -25.7807
16.0298 | -34.7397 | -25.7881
16.0359 | -34.7402 | -25.7957
16.042 | -34.7407 | -25.8037
16.0481 | -34.7412 | -25.8116
16.0542 | -34.7417 | -25.8197
16.0603 | -34.7423 | -25.8275
16.0664 | -34.7428 | -25.8355
16.0725 | -34.7433 | -25.8431
16.0786 | -34.7438 | -25.8513
16.0847 | -34.7443 | -25.859
16.0908 | -34.7448 | -25.8669
16.0969 | -34.7453 | -25.8744
16.103 | -34.7458 | -25.8825
16.1091 | -34.7463 | -25.8902
16.1152 | -34.7468 | -25.8982
16.1214 | -34.7474 | -25.9061
16.1275 | -34.7479 | -25.9137
16.1336 | -34.7484 | -25.9216
16.1397 | -34.7489 | -25.9294
16.1458 | -34.7494 | -25.9372
16.1519 | -34.75 | -25.9452
16.158 | -34.7505 | -25.9531
16.1641 | -34.751 | -25.9609
16.1702 | -34.7515 | -25.9688
16.1763 | -34.752 | -25.9767
16.1824 | -34.7525 | -25.9842
16.1885 | -34.753 | -25.992
16.1946 | -34.7535 | -25.9998
16.2007 | -34.7541 | -26.008
16.2068 | -34.7546 | -26.0157
16.2129 | -34.7551 | -26.0236
16.219 | -34.7556 | -26.0312
16.2251 | -34.7561 | -26.0392
16.2312 | -34.7565 | -26.0469
16.2373 | -34.757 | -26.0547
16.2434 | -34.7575 | -26.0626
16.2495 | -34.7581 | -26.0702
16.2556 | -34.7586 | -26.0782
16.2618 | -34.7591 | -26.0856
16.2679 | -34.7597 | -26.0935
16.274 | -34.7602 | -26.1016
16.2801 | -34.7607 | -26.1092
16.2862 | -34.7613 | -26.117
16.2923 | -34.7617 | -26.1248
16.2984 | -34.7622 | -26.1326
16.3045 | -34.7628 | -26.1403
16.3106 | -34.7633 | -26.1477
16.3167 | -34.7638 | -26.1557
16.3228 | -34.7643 | -26.1635
16.3289 | -34.7648 | -26.1713
16.335 | -34.7653 | -26.1793
16.3411 | -34.7658 | -26.1869
16.3472 | -34.7663 | -26.195
16.3533 | -34.7668 | -26.2028
16.3594 | -34.7673 | -26.2103
16.3655 | -34.7678 | -26.2183
16.3716 | -34.7683 | -26.2258
16.3777 | -34.7688 | -26.2335
16.3838 | -34.7693 | -26.2413
16.3899 | -34.7698 | -26.2492
16.396 | -34.7703 | -26.2571
16.4021 | -34.7708 | -26.2649
16.4083 | -34.7714 | -26.2726
16.4144 | -34.7719 | -26.2802
16.4205 | -34.7724 | -26.2881
16.4266 | -34.7729 | -26.2957
16.4327 | -34.7734 | -26.3032
16.4388 | -34.7739 | -26.3112
16.4449 | -34.7745 | -26.319
16.451 | -34.775 | -26.327
16.4571 | -34.7754 | -26.3349
16.4632 | -34.7759 | -26.3424
16.4693 | -34.7764 | -26.3502
16.4754 | -34.7769 | -26.3582
16.4815 | -34.7774 | -26.3661
16.4876 | -34.778 | -26.3739
16.4937 | -34.7785 | -26.3816
16.4998 | -34.779 | -26.3893
16.5059 | -34.7796 | -26.397
16.512 | -34.7801 | -26.4052
16.5181 | -34.7807 | -26.4132
16.5242 | -34.7811 | -26.4211
16.5303 | -34.7816 | -26.4285
16.5364 | -34.7821 | -26.4363
16.5425 | -34.7826 | -26.4442
16.5487 | -34.7831 | -26.4518
16.5548 | -34.7836 | -26.4594
16.5609 | -34.7841 | -26.4673
16.567 | -34.7846 | -26.4749
16.5731 | -34.7851 | -26.4827
16.5792 | -34.7857 | -26.4904
16.5853 | -34.7862 | -26.498
16.5914 | -34.7867 | -26.5059
16.5975 | -34.7872 | -26.5137
16.6036 | -34.7877 | -26.5213
16.6097 | -34.7882 | -26.529
16.6158 | -34.7887 | -26.5369
16.6219 | -34.7892 | -26.5445
16.628 | -34.7897 | -26.5522
16.6341 | -34.7902 | -26.5596
16.6402 | -34.7908 | -26.5669
16.6463 | -34.7913 | -26.5747
16.6524 | -34.7918 | -26.5824
16.6585 | -34.7923 | -26.5903
16.6646 | -34.7928 | -26.5981
16.6707 | -34.7933 | -26.6056
16.6768 | -34.7938 | -26.6133
16.6829 | -34.7943 | -26.621
16.689 | -34.7948 | -26.6287
16.6952 | -34.7953 | -26.6366
16.7013 | -34.7958 | -26.6444
16.7074 | -34.7964 | -26.6519
16.7135 | -34.7969 | -26.6597
16.7196 | -34.7973 | -26.6674
16.7257 | -34.7978 | -26.6749
16.7318 | -34.7983 | -26.6826
16.7379 | -34.7989 | -26.6905
16.744 | -34.7994 | -26.698
16.7501 | -34.7999 | -26.7057
16.7562 | -34.8004 | -26.7134
16.7623 | -34.8009 | -26.7211
16.7684 | -34.8014 | -26.729
16.7745 | -34.802 | -26.7364
16.7806 | -34.8024 | -26.7441
16.7867 | -34.803 | -26.7517
16.7928 | -34.8035 | -26.7596
16.7989 | -34.8039 | -26.7672
16.805 | -34.8044 | -26.775
16.8111 | -34.805 | -26.7828
16.8172 | -34.8054 | -26.7908
16.8233 | -34.806 | -26.7989
16.8294 | -34.8065 | -26.8068
16.8356 | -34.807 | -26.8148
16.8417 | -34.8075 | -26.8225
16.8478 | -34.808 | -26.8301
16.8539 | -34.8086 | -26.8376
16.86 | -34.8091 | -26.8452
16.8661 | -34.8096 | -26.8531
16.8722 | -34.8102 | -26.8608
16.8783 | -34.8107 | -26.8683
16.8844 | -34.8111 | -26.876
16.8905 | -34.8117 | -26.8838
16.8966 | -34.8122 | -26.8916
16.9027 | -34.8127 | -26.899
16.9088 | -34.8132 | -26.9071
16.9149 | -34.8137 | -26.9151
16.921 | -34.8142 | -26.9228
16.9271 | -34.8148 | -26.9308
16.9332 | -34.8153 | -26.9386
16.9393 | -34.8158 | -26.9467
16.9454 | -34.8162 | -26.9541
16.9515 | -34.8167 | -26.9621
16.9576 | -34.8172 | -26.9696
16.9637 | -34.8178 | -26.9774
16.9698 | -34.8183 | -26.9848
16.9759 | -34.8188 | -26.9924
16.9821 | -34.8193 | -27.0001
16.9882 | -34.8198 | -27.0079
16.9943 | -34.8204 | -27.0157
17.0004 | -34.8209 | -27.0233
17.0065 | -34.8214 | -27.0311
17.0126 | -34.8219 | -27.0387
17.0187 | -34.8224 | -27.0465
17.0248 | -34.8229 | -27.0538
17.0309 | -34.8234 | -27.0614
17.037 | -34.8239 | -27.069
17.0431 | -34.8244 | -27.0764
17.0492 | -34.8249 | -27.0841
17.0553 | -34.8254 | -27.0918
17.0614 | -34.8258 | -27.0994
17.0675 | -34.8264 | -27.1073
17.0736 | -34.8268 | -27.1151
17.0797 | -34.8273 | -27.1225
17.0858 | -34.8278 | -27.1303
17.0919 | -34.8283 | -27.1382
17.098 | -34.8288 | -27.1456
17.1041 | -34.8293 | -27.1531
17.1102 | -34.8298 | -27.1608
17.1163 | -34.8303 | -27.1687
17.1225 | -34.8309 | -27.1761
17.1286 | -34.8314 | -27.1835
17.1347 | -34.8319 | -27.1911
17.1408 | -34.8324 | -27.1988
17.1469 | -34.8329 | -27.2063
17.153 | -34.8334 | -27.2139
17.1591 | -34.8339 | -27.2213
17.1652 | -34.8344 | -27.2287
17.1713 | -34.8349 | -27.2363
17.1774 | -34.8354 | -27.2442
17.1835 | -34.8359 | -27.252
17.1896 | -34.8364 | -27.2594
17.1957 | -34.8369 | -27.2672
17.2018 | -34.8374 | -27.2748
17.2079 | -34.8379 | -27.2824
17.214 | -34.8384 | -27.2899
17.2201 | -34.8388 | -27.2975
17.2262 | -34.8393 | -27.3052
17.2323 | -34.8398 | -27.3127
17.2384 | -34.8404 | -27.3202
17.2445 | -34.8408 | -27.3278
17.2506 | -34.8413 | -27.3356
17.2567 | -34.8418 | -27.3433
17.2628 | -34.8423 | -27.3512
17.269 | -34.8428 | -27.3588
17.2751 | -34.8433 | -27.3663
17.2812 | -34.8438 | -27.374
17.2873 | -34.8443 | -27.3817
17.2934 | -34.8447 | -27.3892
17.2995 | -34.8452 | -27.3966
17.3056 | -34.8457 | -27.404
17.3117 | -34.8462 | -27.4115
17.3178 | -34.8467 | -27.4191
17.3239 | -34.8472 | -27.4267
17.33 | -34.8476 | -27.4341
17.3361 | -34.8481 | -27.4417
17.3422 | -34.8486 | -27.4492
17.3483 | -34.8491 | -27.4568
17.3544 | -34.8496 | -27.4643
17.3605 | -34.8501 | -27.472
17.3666 | -34.8506 | -27.4798
17.3727 | -34.8511 | -27.4873
17.3788 | -34.8516 | -27.4948
17.3849 | -34.8521 | -27.5023
17.391 | -34.8526 | -27.5098
17.3971 | -34.8531 | -27.5174
17.4032 | -34.8536 | -27.5252
17.4094 | -34.8541 | -27.5328
17.4155 | -34.8546 | -27.5404
17.4216 | -34.855 | -27.5479
17.4277 | -34.8555 | -27.5558
17.4338 | -34.8561 | -27.5633
17.4399 | -34.8566 | -27.5709
17.446 | -34.857 | -27.5782
17.4521 | -34.8575 | -27.5857
17.4582 | -34.858 | -27.5932
17.4643 | -34.8585 | -27.6009
17.4704 | -34.859 | -27.6083
17.4765 | -34.8595 | -27.616
17.4826 | -34.86 | -27.6236
17.4887 | -34.8605 | -27.6316
17.4948 | -34.861 | -27.639
17.5009 | -34.8614 | -27.6464
17.507 | -34.8619 | -27.6541
17.5131 | -34.8624 | -27.6617
17.5192 | -34.8629 | -27.669
17.5253 | -34.8633 | -27.6767
17.5314 | -34.8638 | -27.6843
17.5375 | -34.8642 | -27.692
17.5436 | -34.8647 | -27.6998
17.5497 | -34.8652 | -27.7076
17.5559 | -34.8657 | -27.7149
17.562 | -34.8662 | -27.7224
17.5681 | -34.8667 | -27.7299
17.5742 | -34.8671 | -27.7375
17.5803 | -34.8677 | -27.7449
17.5864 | -34.8682 | -27.7522
17.5925 | -34.8687 | -27.76
17.5986 | -34.8692 | -27.7672
17.6047 | -34.8697 | -27.7748
17.6108 | -34.8702 | -27.7823
17.6169 | -34.8707 | -27.7897
17.623 | -34.8712 | -27.7971
17.6291 | -34.8717 | -27.8046
17.6352 | -34.8722 | -27.8119
17.6413 | -34.8727 | -27.8197
17.6474 | -34.8732 | -27.8273
17.6535 | -34.8736 | -27.8349
17.6596 | -34.8741 | -27.8425
17.6657 | -34.8746 | -27.8499
17.6718 | -34.8751 | -27.8576
17.6779 | -34.8755 | -27.865
17.684 | -34.876 | -27.8726
17.6901 | -34.8766 | -27.88
17.6963 | -34.877 | -27.8874
17.7024 | -34.8775 | -27.8948
17.7085 | -34.878 | -27.902
17.7146 | -34.8785 | -27.9093
17.7207 | -34.8789 | -27.9168
17.7268 | -34.8795 | -27.9241
17.7329 | -34.8799 | -27.9318
17.739 | -34.8804 | -27.9396
17.7451 | -34.881 | -27.947
17.7512 | -34.8814 | -27.9546
17.7573 | -34.8819 | -27.9622
17.7634 | -34.8824 | -27.9693
17.7695 | -34.8829 | -27.9768
17.7756 | -34.8834 | -27.9841
17.7817 | -34.8839 | -27.9915
17.7878 | -34.8843 | -27.999
17.7939 | -34.8848 | -28.0063
17.8 | -34.8853 | -28.0137
17.8061 | -34.8858 | -28.0214
17.8122 | -34.8863 | -28.0288
17.8183 | -34.8868 | -28.0365
17.8244 | -34.8873 | -28.0441
17.8305 | -34.8878 | -28.0516
17.8366 | -34.8883 | -28.0587
17.8428 | -34.8888 | -28.0663
17.8489 | -34.8893 | -28.0736
17.855 | -34.8898 | -28.081
17.8611 | -34.8903 | -28.0882
17.8672 | -34.8908 | -28.0954
17.8733 | -34.8913 | -28.1031
17.8794 | -34.8917 | -28.1108
17.8855 | -34.8923 | -28.1182
17.8916 | -34.8927 | -28.1254
17.8977 | -34.8932 | -28.1329
17.9038 | -34.8937 | -28.1402
17.9099 | -34.8941 | -28.1477
17.916 | -34.8946 | -28.1551
17.9221 | -34.8951 | -28.1625
17.9282 | -34.8956 | -28.1699
17.9343 | -34.8961 | -28.1774
17.9404 | -34.8966 | -28.1847
17.9465 | -34.897 | -28.1922
17.9526 | -34.8975 | -28.1999
17.9587 | -34.898 | -28.2074
17.9648 | -34.8985 | -28.2149
17.9709 | -34.899 | -28.2222
17.977 | -34.8994 | -28.2294
17.9832 | -34.8999 | -28.2372
17.9893 | -34.9004 | -28.2447
17.9954 | -34.9009 | -28.2523
18.0015 | -34.9015 | -28.2597
18.0076 | -34.9019 | -28.2667
18.0137 | -34.9023 | -28.2746
18.0198 | -34.9028 | -28.282
18.0259 | -34.9032 | -28.2895
18.032 | -34.9037 | -28.2966
18.0381 | -34.9042 | -28.3038
18.0442 | -34.9047 | -28.3113
18.0503 | -34.9052 | -28.3188
18.0564 | -34.9056 | -28.3261
18.0625 | -34.9061 | -28.3337
18.0686 | -34.9066 | -28.3409
18.0747 | -34.9071 | -28.3483
18.0808 | -34.9076 | -28.3555
18.0869 | -34.9081 | -28.363
18.093 | -34.9086 | -28.3705
18.0991 | -34.909 | -28.3782
18.1052 | -34.9095 | -28.386
18.1113 | -34.91 | -28.3932
18.1174 | -34.9105 | -28.4006
18.1236 | -34.911 | -28.4081
18.1297 | -34.9115 | -28.4156
18.1358 | -34.912 | -28.4231
18.1419 | -34.9125 | -28.4307
18.148 | -34.9131 | -28.4381
18.1541 | -34.9136 | -28.4453
18.1602 | -34.9141 | -28.453
18.1663 | -34.9146 | -28.4603
18.1724 | -34.9151 | -28.468
18.1785 | -34.9156 | -28.4756
18.1846 | -34.916 | -28.4828
18.1907 | -34.9165 | -28.4906
18.1968 | -34.9169 | -28.4982
18.2029 | -34.9174 | -28.5052
18.209 | -34.9179 | -28.5129
18.2151 | -34.9184 | -28.5206
18.2212 | -34.9189 | -28.5277
18.2273 | -34.9194 | -28.5352
18.2334 | -34.9199 | -28.5424
18.2395 | -34.9204 | -28.5497
18.2456 | -34.9208 | -28.5569
18.2517 | -34.9213 | -28.5643
18.2578 | -34.9218 | -28.5715
18.2639 | -34.9223 | -28.5791
18.2701 | -34.9228 | -28.5865
18.2762 | -34.9233 | -28.5941
18.2823 | -34.9238 | -28.6015
18.2884 | -34.9243 | -28.6089
18.2945 | -34.9248 | -28.6161
18.3006 | -34.9253 | -28.6236
18.3067 | -34.9258 | -28.6309
18.3128 | -34.9263 | -28.6382
18.3189 | -34.9268 | -28.6458
18.325 | -34.9273 | -28.6528
18.3311 | -34.9278 | -28.6604
18.3372 | -34.9282 | -28.6677
18.3433 | -34.9287 | -28.6748
18.3494 | -34.9292 | -28.6823
18.3555 | -34.9296 | -28.6892
18.3616 | -34.9301 | -28.6963
18.3677 | -34.9307 | -28.7035
18.3738 | -34.9311 | -28.7108
18.3799 | -34.9316 | -28.7179
18.386 | -34.9321 | -28.7252
18.3921 | -34.9326 | -28.7325
18.3982 | -34.9331 | -28.7397
18.4043 | -34.9335 | -28.7469
18.4105 | -34.934 | -28.7544
18.4166 | -34.9345 | -28.7618
18.4227 | -34.935 | -28.7692
18.4288 | -34.9355 | -28.7765
18.4349 | -34.936 | -28.7837
18.441 | -34.9365 | -28.7909
18.4471 | -34.937 | -28.7982
18.4532 | -34.9375 | -28.8055
18.4593 | -34.9379 | -28.8126
18.4654 | -34.9383 | -28.8197
18.4715 | -34.9388 | -28.827
18.4776 | -34.9393 | -28.8344
18.4837 | -34.9398 | -28.8418
18.4898 | -34.9403 | -28.8493
18.4959 | -34.9408 | -28.8569
18.502 | -34.9414 | -28.8643
18.5081 | -34.9419 | -28.8719
18.5142 | -34.9423 | -28.8795
18.5203 | -34.9428 | -28.8866
18.5264 | -34.9433 | -28.8938
18.5325 | -34.9438 | -28.9013
18.5386 | -34.9443 | -28.9089
18.5447 | -34.9448 | -28.916
18.5508 | -34.9453 | -28.9231
18.557 | -34.9458 | -28.9307
18.5631 | -34.9462 | -28.9382
18.5692 | -34.9467 | -28.9455
18.5753 | -34.9472 | -28.9525
18.5814 | -34.9477 | -28.9601
18.5875 | -34.9482 | -28.9676
18.5936 | -34.9487 | -28.9753
18.5997 | -34.9492 | -28.9827
18.6058 | -34.9496 | -28.9901
18.6119 | -34.9501 | -28.9975
18.618 | -34.9506 | -29.005
18.6241 | -34.9511 | -29.0124
18.6302 | -34.9517 | -29.0199
18.6363 | -34.9521 | -29.027
18.6424 | -34.9526 | -29.0342
18.6485 | -34.9531 | -29.0417
18.6546 | -34.9536 | -29.0487
18.6607 | -34.954 | -29.056
18.6668 | -34.9545 | -29.0634
18.6729 | -34.955 | -29.0707
18.679 | -34.9555 | -29.0778
18.6851 | -34.956 | -29.0851
18.6912 | -34.9565 | -29.0924
18.6974 | -34.957 | -29.0996
18.7035 | -34.9575 | -29.1068
18.7096 | -34.958 | -29.1141
18.7157 | -34.9585 | -29.1213
18.7218 | -34.959 | -29.1286
18.7279 | -34.9594 | -29.1357
18.734 | -34.9599 | -29.1431
18.7401 | -34.9604 | -29.1501
18.7462 | -34.9609 | -29.1568
18.7523 | -34.9614 | -29.1641
18.7584 | -34.9619 | -29.1715
18.7645 | -34.9625 | -29.1786
18.7706 | -34.963 | -29.1859
18.7767 | -34.9635 | -29.193
18.7828 | -34.964 | -29.1999
18.7889 | -34.9644 | -29.2072
18.795 | -34.9649 | -29.2145
18.8011 | -34.9654 | -29.2218
18.8072 | -34.9659 | -29.229
18.8133 | -34.9664 | -29.2363
18.8194 | -34.9668 | -29.2434
18.8255 | -34.9673 | -29.2505
18.8316 | -34.9678 | -29.2581
18.8377 | -34.9682 | -29.2653
18.8439 | -34.9688 | -29.2725
18.85 | -34.9692 | -29.2799
18.8561 | -34.9697 | -29.2872
18.8622 | -34.9702 | -29.2946
18.8683 | -34.9707 | -29.3016
18.8744 | -34.9712 | -29.3088
18.8805 | -34.9717 | -29.3162
18.8866 | -34.9722 | -29.3235
18.8927 | -34.9727 | -29.3305
18.8988 | -34.9731 | -29.3378
18.9049 | -34.9736 | -29.3451
18.911 | -34.9741 | -29.3523
18.9171 | -34.9747 | -29.3595
18.9232 | -34.9752 | -29.3665
18.9293 | -34.9756 | -29.3738
18.9354 | -34.9761 | -29.3809
18.9415 | -34.9766 | -29.3882
18.9476 | -34.9771 | -29.3955
18.9537 | -34.9776 | -29.403
18.9598 | -34.9781 | -29.4096
18.9659 | -34.9786 | -29.4168
18.972 | -34.9791 | -29.4241
18.9781 | -34.9795 | -29.4313
18.9843 | -34.98 | -29.4385
18.9904 | -34.9805 | -29.4457
18.9965 | -34.981 | -29.4527
19.0026 | -34.9815 | -29.4601
19.0087 | -34.982 | -29.4673
19.0148 | -34.9825 | -29.4743
19.0209 | -34.983 | -29.4818
19.027 | -34.9835 | -29.4892
19.0331 | -34.984 | -29.4965
19.0392 | -34.9845 | -29.5038
19.0453 | -34.985 | -29.5111
19.0514 | -34.9854 | -29.5183
19.0575 | -34.9859 | -29.5255
19.0636 | -34.9864 | -29.5323
19.0697 | -34.9869 | -29.5398
19.0758 | -34.9874 | -29.5468
19.0819 | -34.9879 | -29.5543
19.088 | -34.9884 | -29.5612
19.0941 | -34.9889 | -29.568
19.1002 | -34.9893 | -29.575
19.1063 | -34.9899 | -29.5822
19.1124 | -34.9904 | -29.5893
19.1185 | -34.9908 | -29.5966
19.1246 | -34.9913 | -29.6036
19.1308 | -34.9918 | -29.6105
19.1369 | -34.9923 | -29.6175
19.143 | -34.9928 | -29.6247
19.1491 | -34.9933 | -29.6316
19.1552 | -34.9938 | -29.6392
19.1613 | -34.9943 | -29.6464
19.1674 | -34.9947 | -29.6536
19.1735 | -34.9952 | -29.6609
19.1796 | -34.9957 | -29.6678
19.1857 | -34.9962 | -29.6752
19.1918 | -34.9967 | -29.6825
19.1979 | -34.9971 | -29.6896
19.204 | -34.9976 | -29.6966
19.2101 | -34.9981 | -29.7038
19.2162 | -34.9986 | -29.7112
19.2223 | -34.9991 | -29.7179
19.2284 | -34.9996 | -29.7254
19.2345 | -35.0001 | -29.7327
19.2406 | -35.0006 | -29.74
19.2467 | -35.0011 | -29.7471
19.2528 | -35.0016 | -29.7542
19.2589 | -35.002 | -29.7613
19.265 | -35.0025 | -29.7686
19.2712 | -35.003 | -29.7758
19.2773 | -35.0035 | -29.783
19.2834 | -35.004 | -29.79
19.2895 | -35.0045 | -29.7974
19.2956 | -35.005 | -29.8048
19.3017 | -35.0055 | -29.8118
19.3078 | -35.006 | -29.819
19.3139 | -35.0064 | -29.8262
19.32 | -35.0069 | -29.8332
19.3261 | -35.0074 | -29.8403
19.3322 | -35.0079 | -29.8475
19.3383 | -35.0084 | -29.8544
19.3444 | -35.0089 | -29.8615
19.3505 | -35.0094 | -29.8686
19.3566 | -35.0098 | -29.8756
19.3627 | -35.0103 | -29.8829
19.3688 | -35.0109 | -29.8901
19.3749 | -35.0114 | -29.8975
19.381 | -35.0118 | -29.9043
19.3871 | -35.0123 | -29.9115
19.3932 | -35.0128 | -29.9189
19.3993 | -35.0132 | -29.926
19.4054 | -35.0138 | -29.9334
19.4115 | -35.0143 | -29.9408
19.4177 | -35.0147 | -29.9476
19.4238 | -35.0153 | -29.9545
19.4299 | -35.0158 | -29.9617
19.436 | -35.0163 | -29.9692
19.4421 | -35.0168 | -29.9761
19.4482 | -35.0173 | -29.9833
19.4543 | -35.0178 | -29.9902
19.4604 | -35.0183 | -29.9972
19.4665 | -35.0187 | -30.0042
19.4726 | -35.0193 | -30.011
19.4787 | -35.0197 | -30.0182
19.4848 | -35.0202 | -30.0252
19.4909 | -35.0207 | -30.0321
19.497 | -35.0212 | -30.0391
19.5031 | -35.0216 | -30.0461
19.5092 | -35.0221 | -30.0532
19.5153 | -35.0227 | -30.0605
19.5214 | -35.0232 | -30.0673
19.5275 | -35.0236 | -30.0744
19.5336 | -35.0241 | -30.0815
19.5397 | -35.0246 | -30.0883
19.5458 | -35.0251 | -30.0955
19.5519 | -35.0256 | -30.1025
19.5581 | -35.0261 | -30.1092
19.5642 | -35.0266 | -30.1161
19.5703 | -35.027 | -30.123
19.5764 | -35.0275 | -30.1302
19.5825 | -35.028 | -30.1371
19.5886 | -35.0285 | -30.1441
19.5947 | -35.029 | -30.1512
19.6008 | -35.0294 | -30.1584
19.6069 | -35.0299 | -30.1653
19.613 | -35.0304 | -30.1727
19.6191 | -35.0309 | -30.1798
19.6252 | -35.0314 | -30.1868
19.6313 | -35.0318 | -30.1936
19.6374 | -35.0323 | -30.2008
19.6435 | -35.0328 | -30.2077
19.6496 | -35.0333 | -30.2147
19.6557 | -35.0338 | -30.2218
19.6618 | -35.0343 | -30.2285
19.6679 | -35.0347 | -30.2359
19.674 | -35.0352 | -30.2431
19.6801 | -35.0357 | -30.25
19.6862 | -35.0362 | -30.2569
19.6923 | -35.0367 | -30.2639
19.6984 | -35.0372 | -30.2706
19.7046 | -35.0377 | -30.2779
19.7107 | -35.0382 | -30.2849
19.7168 | -35.0387 | -30.2921
19.7229 | -35.0391 | -30.2991
19.729 | -35.0396 | -30.3062
19.7351 | -35.0401 | -30.3128
19.7412 | -35.0406 | -30.3199
19.7473 | -35.0411 | -30.3269
19.7534 | -35.0416 | -30.334
19.7595 | -35.0421 | -30.3413
19.7656 | -35.0426 | -30.3484
19.7717 | -35.043 | -30.3553
19.7778 | -35.0435 | -30.3625
19.7839 | -35.044 | -30.3692
19.79 | -35.0444 | -30.3762
19.7961 | -35.0449 | -30.3832
19.8022 | -35.0453 | -30.3902
19.8083 | -35.0459 | -30.3972
19.8144 | -35.0463 | -30.4041
19.8205 | -35.0468 | -30.4113
19.8266 | -35.0473 | -30.4182
19.8327 | -35.0477 | -30.4254
19.8388 | -35.0482 | -30.4326
19.845 | -35.0487 | -30.4394
19.8511 | -35.0492 | -30.4466
19.8572 | -35.0497 | -30.4537
19.8633 | -35.0502 | -30.4606
19.8694 | -35.0507 | -30.4678
19.8755 | -35.0512 | -30.4748
19.8816 | -35.0517 | -30.4819
19.8877 | -35.0521 | -30.4889
19.8938 | -35.0526 | -30.4958
19.8999 | -35.0531 | -30.5029
19.906 | -35.0536 | -30.51
19.9121 | -35.0541 | -30.5169
19.9182 | -35.0546 | -30.5238
19.9243 | -35.0551 | -30.531
19.9304 | -35.0555 | -30.5384
19.9365 | -35.056 | -30.5455
19.9426 | -35.0565 | -30.5522
19.9487 | -35.057 | -30.559
19.9548 | -35.0575 | -30.5662
19.9609 | -35.058 | -30.5731
19.967 | -35.0585 | -30.5799
19.9731 | -35.059 | -30.5869
19.9792 | -35.0594 | -30.5935
19.9853 | -35.0599 | -30.6007
19.9915 | -35.0605 | -30.608
19.9976 | -35.0609 | -30.6148
20.0037 | -35.0614 | -30.6218
20.0098 | -35.0618 | -30.6287
20.0159 | -35.0623 | -30.6355
20.022 | -35.0628 | -30.6428
20.0281 | -35.0633 | -30.6499
20.0342 | -35.0637 | -30.6566
20.0403 | -35.0642 | -30.6632
20.0464 | -35.0647 | -30.6706
20.0525 | -35.0652 | -30.6773
20.0586 | -35.0657 | -30.6842
20.0647 | -35.0663 | -30.6915
20.0708 | -35.0668 | -30.6987
20.0769 | -35.0672 | -30.7057
20.083 | -35.0677 | -30.7121
20.0891 | -35.0681 | -30.7191
20.0952 | -35.0687 | -30.7259
20.1013 | -35.0691 | -30.7327
20.1074 | -35.0695 | -30.74
20.1135 | -35.07 | -30.7468
20.1196 | -35.0705 | -30.7535
20.1257 | -35.071 | -30.7605
20.1319 | -35.0715 | -30.7673
20.138 | -35.072 | -30.7741
20.1441 | -35.0725 | -30.7809
20.1502 | -35.0729 | -30.788
20.1563 | -35.0734 | -30.7946
20.1624 | -35.0738 | -30.8014
20.1685 | -35.0743 | -30.8085
20.1746 | -35.0748 | -30.8155
20.1807 | -35.0752 | -30.8225
20.1868 | -35.0757 | -30.8294
20.1929 | -35.0762 | -30.836
20.199 | -35.0766 | -30.8431
20.2051 | -35.0771 | -30.8501
20.2112 | -35.0776 | -30.8568
20.2173 | -35.078 | -30.8638
20.2234 | -35.0785 | -30.8707
20.2295 | -35.0789 | -30.8775
20.2356 | -35.0794 | -30.8846
20.2417 | -35.0799 | -30.8915
20.2478 | -35.0803 | -30.8984
20.2539 | -35.0807 | -30.9055
20.26 | -35.0812 | -30.9122
20.2661 | -35.0817 | -30.9194
20.2723 | -35.0821 | -30.9264
20.2784 | -35.0826 | -30.9334
20.2845 | -35.0831 | -30.9402
20.2906 | -35.0836 | -30.9472
20.2967 | -35.0841 | -30.9536
20.3028 | -35.0845 | -30.9603
20.3089 | -35.085 | -30.9672
20.315 | -35.0855 | -30.9742
20.3211 | -35.0859 | -30.981
20.3272 | -35.0865 | -30.9878
20.3333 | -35.0869 | -30.9946
20.3394 | -35.0873 | -31.0015
20.3455 | -35.0878 | -31.0083
20.3516 | -35.0883 | -31.0154
20.3577 | -35.0888 | -31.0221
20.3638 | -35.0892 | -31.0291
20.3699 | -35.0897 | -31.0355
20.376 | -35.0902 | -31.0424
20.3821 | -35.0907 | -31.0492
20.3882 | -35.0912 | -31.0561
20.3943 | -35.0917 | -31.0632
20.4004 | -35.0922 | -31.0699
20.4065 | -35.0926 | -31.0766
20.4126 | -35.0931 | -31.0833
20.4188 | -35.0936 | -31.0905
20.4249 | -35.094 | -31.0973
20.431 | -35.0945 | -31.1043
20.4371 | -35.095 | -31.1105
20.4432 | -35.0954 | -31.1171
20.4493 | -35.0959 | -31.1239
20.4554 | -35.0964 | -31.1308
20.4615 | -35.0968 | -31.1378
20.4676 | -35.0973 | -31.1448
20.4737 | -35.0978 | -31.1512
20.4798 | -35.0983 | -31.1581
20.4859 | -35.0987 | -31.1648
20.492 | -35.0992 | -31.1715
20.4981 | -35.0997 | -31.178
20.5042 | -35.1002 | -31.1847
20.5103 | -35.1006 | -31.1914
20.5164 | -35.1011 | -31.1981
20.5225 | -35.1015 | -31.205
20.5286 | -35.1019 | -31.212
20.5347 | -35.1024 | -31.2187
20.5408 | -35.1029 | -31.2255
20.5469 | -35.1034 | -31.232
20.553 | -35.1039 | -31.2387
20.5592 | -35.1043 | -31.2455
20.5653 | -35.1047 | -31.2524
20.5714 | -35.1052 | -31.2588
20.5775 | -35.1056 | -31.2657
20.5836 | -35.1061 | -31.2729
20.5897 | -35.1066 | -31.2795
20.5958 | -35.107 | -31.2863
20.6019 | -35.1075 | -31.293
20.608 | -35.1079 | -31.2999
20.6141 | -35.1084 | -31.3064
20.6202 | -35.1088 | -31.3135
20.6263 | -35.1093 | -31.32
20.6324 | -35.1097 | -31.3267
20.6385 | -35.1102 | -31.3332
20.6446 | -35.1107 | -31.34
20.6507 | -35.1111 | -31.3463
20.6568 | -35.1116 | -31.3529
20.6629 | -35.112 | -31.3597
20.669 | -35.1125 | -31.3665
20.6751 | -35.113 | -31.3733
20.6812 | -35.1134 | -31.3799
20.6873 | -35.1139 | -31.3868
20.6934 | -35.1144 | -31.3934
20.6995 | -35.1148 | -31.4002
20.7057 | -35.1153 | -31.4069
20.7118 | -35.1158 | -31.4133
20.7179 | -35.1163 | -31.4197
20.724 | -35.1167 | -31.4263
20.7301 | -35.1171 | -31.4331
20.7362 | -35.1175 | -31.44
20.7423 | -35.1181 | -31.4461
20.7484 | -35.1185 | -31.4529
20.7545 | -35.119 | -31.4595
20.7606 | -35.1195 | -31.4662
20.7667 | -35.12 | -31.4728
20.7728 | -35.1205 | -31.4796
20.7789 | -35.1209 | -31.4863
20.785 | -35.1214 | -31.4931
20.7911 | -35.1218 | -31.5001
20.7972 | -35.1222 | -31.5069
20.8033 | -35.1227 | -31.5137
20.8094 | -35.1232 | -31.5205
20.8155 | -35.1236 | -31.5273
20.8216 | -35.124 | -31.534
20.8277 | -35.1244 | -31.5406
20.8338 | -35.1249 | -31.5476
20.8399 | -35.1254 | -31.5543
20.8461 | -35.1258 | -31.5611
20.8522 | -35.1263 | -31.5678
20.8583 | -35.1268 | -31.5742
20.8644 | -35.1273 | -31.581
20.8705 | -35.1278 | -31.5879
20.8766 | -35.1282 | -31.5947
20.8827 | -35.1288 | -31.6015
20.8888 | -35.1292 | -31.6081
20.8949 | -35.1296 | -31.6148
20.901 | -35.13 | -31.6215
20.9071 | -35.1305 | -31.6279
20.9132 | -35.1309 | -31.6346
20.9193 | -35.1314 | -31.6413
20.9254 | -35.1318 | -31.6481
20.9315 | -35.1323 | -31.6545
20.9376 | -35.1327 | -31.6612
20.9437 | -35.1331 | -31.668
20.9498 | -35.1336 | -31.6742
20.9559 | -35.1341 | -31.6805
20.962 | -35.1346 | -31.6873
20.9681 | -35.1351 | -31.6941
20.9742 | -35.1356 | -31.7009
20.9803 | -35.136 | -31.7077
20.9864 | -35.1365 | -31.7145
20.9926 | -35.137 | -31.7213
20.9987 | -35.1374 | -31.7277
21.0048 | -35.1379 | -31.7346
21.0109 | -35.1383 | -31.7413
21.017 | -35.1388 | -31.7478
21.0231 | -35.1392 | -31.7544
21.0292 | -35.1397 | -31.761
21.0353 | -35.1401 | -31.7676
21.0414 | -35.1405 | -31.7743
21.0475 | -35.141 | -31.7815
21.0536 | -35.1415 | -31.7879
21.0597 | -35.142 | -31.7944
21.0658 | -35.1425 | -31.8012
21.0719 | -35.143 | -31.8081
21.078 | -35.1434 | -31.8145
21.0841 | -35.1439 | -31.8212
21.0902 | -35.1444 | -31.8279
21.0963 | -35.1448 | -31.8345
21.1024 | -35.1453 | -31.841
21.1085 | -35.1457 | -31.8477
21.1146 | -35.1461 | -31.8545
21.1207 | -35.1465 | -31.8608
21.1268 | -35.147 | -31.8672
21.133 | -35.1475 | -31.8738
21.1391 | -35.148 | -31.8805
21.1452 | -35.1485 | -31.8872
21.1513 | -35.1489 | -31.8938
21.1574 | -35.1493 | -31.9007
21.1635 | -35.1498 | -31.907
21.1696 | -35.1503 | -31.9133
21.1757 | -35.1507 | -31.9194
21.1818 | -35.1512 | -31.926
21.1879 | -35.1517 | -31.9322
21.194 | -35.1521 | -31.9387
21.2001 | -35.1526 | -31.9452
21.2062 | -35.153 | -31.952
21.2123 | -35.1535 | -31.9587
21.2184 | -35.1539 | -31.9655
21.2245 | -35.1545 | -31.9721
21.2306 | -35.1549 | -31.9784
21.2367 | -35.1554 | -31.985
21.2428 | -35.1559 | -31.9909
21.2489 | -35.1564 | -31.9976
21.255 | -35.1568 | -32.0043
21.2611 | -35.1572 | -32.011
21.2672 | -35.1577 | -32.0175
21.2733 | -35.1581 | -32.0242
21.2795 | -35.1586 | -32.0303
21.2856 | -35.1591 | -32.037
21.2917 | -35.1596 | -32.0437
21.2978 | -35.16 | -32.0506
21.3039 | -35.1605 | -32.0572
21.31 | -35.1609 | -32.0633
21.3161 | -35.1615 | -32.0693
21.3222 | -35.1619 | -32.0756
21.3283 | -35.1624 | -32.0817
21.3344 | -35.1628 | -32.0884
21.3405 | -35.1633 | -32.095
21.3466 | -35.1637 | -32.1015
21.3527 | -35.1641 | -32.1079
21.3588 | -35.1647 | -32.1142
21.3649 | -35.1651 | -32.1207
21.371 | -35.1656 | -32.1272
21.3771 | -35.166 | -32.1336
21.3832 | -35.1664 | -32.1396
21.3893 | -35.167 | -32.1462
21.3954 | -35.1673 | -32.1526
21.4015 | -35.1678 | -32.1589
21.4076 | -35.1683 | -32.1653
21.4137 | -35.1687 | -32.1717
21.4199 | -35.1692 | -32.1783
21.426 | -35.1696 | -32.1849
21.4321 | -35.1701 | -32.1917
21.4382 | -35.1706 | -32.1981
21.4443 | -35.1711 | -32.2042
21.4504 | -35.1715 | -32.211
21.4565 | -35.1719 | -32.2173
21.4626 | -35.1725 | -32.2239
21.4687 | -35.1729 | -32.2306
21.4748 | -35.1733 | -32.2374
21.4809 | -35.1738 | -32.2439
21.487 | -35.1742 | -32.2503
21.4931 | -35.1747 | -32.2561
21.4992 | -35.1752 | -32.2621
21.5053 | -35.1756 | -32.2685
21.5114 | -35.1761 | -32.2746
21.5175 | -35.1765 | -32.281
21.5236 | -35.1769 | -32.2874
21.5297 | -35.1774 | -32.2938
21.5358 | -35.1778 | -32.3004
21.5419 | -35.1782 | -32.3072
21.548 | -35.1787 | -32.3137
21.5541 | -35.1791 | -32.3203
21.5602 | -35.1795 | -32.3271
21.5664 | -35.1799 | -32.3331
21.5725 | -35.1804 | -32.3396
21.5786 | -35.1809 | -32.3459
21.5847 | -35.1812 | -32.3524
21.5908 | -35.1817 | -32.359
21.5969 | -35.1822 | -32.365
21.603 | -35.1826 | -32.3714
21.6091 | -35.1831 | -32.3779
21.6152 | -35.1836 | -32.3846
21.6213 | -35.1841 | -32.3911
21.6274 | -35.1845 | -32.3976
21.6335 | -35.1849 | -32.4039
21.6396 | -35.1854 | -32.4106
21.6457 | -35.1858 | -32.4175
21.6518 | -35.1863 | -32.4238
21.6579 | -35.1868 | -32.4299
21.664 | -35.1872 | -32.4364
21.6701 | -35.1877 | -32.4427
21.6762 | -35.1881 | -32.4491
21.6823 | -35.1886 | -32.4555
21.6884 | -35.189 | -32.4621
21.6945 | -35.1894 | -32.4686
21.7006 | -35.1898 | -32.4752
21.7068 | -35.1903 | -32.4818
21.7129 | -35.1907 | -32.488
21.719 | -35.1911 | -32.4946
21.7251 | -35.1916 | -32.501
21.7312 | -35.1921 | -32.5074
21.7373 | -35.1925 | -32.5135
21.7434 | -35.193 | -32.5201
21.7495 | -35.1934 | -32.5267
21.7556 | -35.1939 | -32.5334
21.7617 | -35.1944 | -32.5399
21.7678 | -35.1949 | -32.5467
21.7739 | -35.1953 | -32.553
21.78 | -35.1958 | -32.5592
21.7861 | -35.1962 | -32.5657
21.7922 | -35.1967 | -32.5723
21.7983 | -35.1971 | -32.5787
21.8044 | -35.1976 | -32.5852
21.8105 | -35.198 | -32.5915
21.8166 | -35.1985 | -32.598
21.8227 | -35.1989 | -32.6042
21.8288 | -35.1994 | -32.6104
21.8349 | -35.1998 | -32.6166
21.841 | -35.2002 | -32.6234
21.8471 | -35.2006 | -32.6297
21.8533 | -35.2011 | -32.6359
21.8594 | -35.2015 | -32.6426
21.8655 | -35.202 | -32.6495
21.8716 | -35.2025 | -32.6558
21.8777 | -35.2029 | -32.6621
21.8838 | -35.2033 | -32.6682
21.8899 | -35.2038 | -32.6745
21.896 | -35.2042 | -32.681
21.9021 | -35.2046 | -32.6873
21.9082 | -35.2051 | -32.6939
21.9143 | -35.2055 | -32.7003
21.9204 | -35.206 | -32.7066
21.9265 | -35.2064 | -32.713
21.9326 | -35.2068 | -32.7192
21.9387 | -35.2073 | -32.7256
21.9448 | -35.2077 | -32.7319
21.9509 | -35.2082 | -32.7379
21.957 | -35.2087 | -32.7441
21.9631 | -35.2091 | -32.7506
21.9692 | -35.2095 | -32.7566
21.9753 | -35.2099 | -32.7631
21.9814 | -35.2104 | -32.7693
21.9875 | -35.2108 | -32.7755
21.9937 | -35.2112 | -32.7821
21.9998 | -35.2117 | -32.7884
22.0059 | -35.2121 | -32.7947
22.012 | -35.2126 | -32.8011
22.0181 | -35.213 | -32.8075
22.0242 | -35.2135 | -32.8138
22.0303 | -35.2139 | -32.8203
22.0364 | -35.2143 | -32.8265
22.0425 | -35.2148 | -32.8325
22.0486 | -35.2152 | -32.8388
22.0547 | -35.2157 | -32.8448
22.0608 | -35.2161 | -32.851
22.0669 | -35.2165 | -32.8573
22.073 | -35.217 | -32.8634
22.0791 | -35.2174 | -32.8698
22.0852 | -35.2179 | -32.8763
22.0913 | -35.2183 | -32.8828
22.0974 | -35.2187 | -32.8886
22.1035 | -35.2191 | -32.8949
22.1096 | -35.2196 | -32.9014
22.1157 | -35.22 | -32.9077
22.1218 | -35.2204 | -32.9142
22.1279 | -35.2209 | -32.9202
22.134 | -35.2213 | -32.9263
22.1402 | -35.2218 | -32.9326
22.1463 | -35.2222 | -32.9387
22.1524 | -35.2227 | -32.9452
22.1585 | -35.2231 | -32.9518
22.1646 | -35.2235 | -32.958
22.1707 | -35.2239 | -32.9645
22.1768 | -35.2243 | -32.9708
22.1829 | -35.2248 | -32.9771
22.189 | -35.2252 | -32.9836
22.1951 | -35.2256 | -32.9898
22.2012 | -35.2261 | -32.9962
22.2073 | -35.2265 | -33.0023
22.2134 | -35.2269 | -33.0087
22.2195 | -35.2274 | -33.0149
22.2256 | -35.2278 | -33.0213
22.2317 | -35.2282 | -33.0276
22.2378 | -35.2287 | -33.0339
22.2439 | -35.2291 | -33.0402
22.25 | -35.2295 | -33.0466
22.2561 | -35.23 | -33.0529
22.2622 | -35.2304 | -33.0594
22.2683 | -35.2308 | -33.0655
22.2744 | -35.2313 | -33.0717
22.2806 | -35.2317 | -33.078
22.2867 | -35.2321 | -33.0844
22.2928 | -35.2326 | -33.0908
22.2989 | -35.233 | -33.0972
22.305 | -35.2335 | -33.1033
22.3111 | -35.2339 | -33.1096
22.3172 | -35.2344 | -33.116
22.3233 | -35.2348 | -33.1223
22.3294 | -35.2353 | -33.1285
22.3355 | -35.2357 | -33.1343
22.3416 | -35.2361 | -33.1408
22.3477 | -35.2366 | -33.1471
22.3538 | -35.237 | -33.1533
22.3599 | -35.2374 | -33.1597
22.366 | -35.2379 | -33.1661
22.3721 | -35.2383 | -33.1719
22.3782 | -35.2388 | -33.1782
22.3843 | -35.2392 | -33.1846
22.3904 | -35.2397 | -33.1912
22.3965 | -35.2401 | -33.1976
22.4026 | -35.2405 | -33.2038
22.4087 | -35.241 | -33.2101
22.4148 | -35.2414 | -33.2164
22.4209 | -35.2418 | -33.2228
22.4271 | -35.2423 | -33.2288
22.4332 | -35.2427 | -33.2348
22.4393 | -35.2432 | -33.2412
22.4454 | -35.2436 | -33.247
22.4515 | -35.244 | -33.2532
22.4576 | -35.2445 | -33.2595
22.4637 | -35.2449 | -33.2659
22.4698 | -35.2454 | -33.2721
22.4759 | -35.2458 | -33.278
22.482 | -35.2462 | -33.2843
22.4881 | -35.2467 | -33.2907
22.4942 | -35.2471 | -33.2971
22.5003 | -35.2475 | -33.3037
22.5064 | -35.2479 | -33.3097
22.5125 | -35.2483 | -33.3159
22.5186 | -35.2487 | -33.3219
22.5247 | -35.2492 | -33.3281
22.5308 | -35.2496 | -33.3347
22.5369 | -35.2501 | -33.3411
22.543 | -35.2505 | -33.3473
22.5491 | -35.251 | -33.353
22.5552 | -35.2514 | -33.3596
22.5613 | -35.2519 | -33.3663
22.5675 | -35.2523 | -33.3727
22.5736 | -35.2527 | -33.3793
22.5797 | -35.2532 | -33.3854
22.5858 | -35.2536 | -33.3915
22.5919 | -35.254 | -33.3976
22.598 | -35.2545 | -33.4039
22.6041 | -35.2549 | -33.41
22.6102 | -35.2553 | -33.4163
22.6163 | -35.2556 | -33.4228
22.6224 | -35.2561 | -33.4294
22.6285 | -35.2565 | -33.4355
22.6346 | -35.257 | -33.4416
22.6407 | -35.2574 | -33.4479
22.6468 | -35.2578 | -33.4542
22.6529 | -35.2583 | -33.4603
22.659 | -35.2588 | -33.4665
22.6651 | -35.2592 | -33.4728
22.6712 | -35.2596 | -33.4788
22.6773 | -35.2601 | -33.485
22.6834 | -35.2605 | -33.4913
22.6895 | -35.261 | -33.4973
22.6956 | -35.2614 | -33.5034
22.7017 | -35.2618 | -33.5096
22.7079 | -35.2622 | -33.516
22.714 | -35.2627 | -33.5223
22.7201 | -35.2631 | -33.5282
22.7262 | -35.2635 | -33.5344
22.7323 | -35.264 | -33.5407
22.7384 | -35.2644 | -33.5468
22.7445 | -35.2648 | -33.5532
22.7506 | -35.2652 | -33.5599
22.7567 | -35.2657 | -33.5662
22.7628 | -35.2661 | -33.5723
22.7689 | -35.2665 | -33.5784
22.775 | -35.2669 | -33.5846
22.7811 | -35.2674 | -33.591
22.7872 | -35.2677 | -33.597
22.7933 | -35.2682 | -33.6029
22.7994 | -35.2686 | -33.609
22.8055 | -35.269 | -33.6153
22.8116 | -35.2695 | -33.6213
22.8177 | -35.2699 | -33.6276
22.8238 | -35.2703 | -33.6337
22.8299 | -35.2707 | -33.64
22.836 | -35.2711 | -33.6462
22.8421 | -35.2716 | -33.6524
22.8482 | -35.272 | -33.6587
22.8544 | -35.2724 | -33.6648
22.8605 | -35.2729 | -33.6711
22.8666 | -35.2733 | -33.6774
22.8727 | -35.2738 | -33.6835
22.8788 | -35.2742 | -33.6895
22.8849 | -35.2746 | -33.6957
22.891 | -35.2751 | -33.7021
22.8971 | -35.2755 | -33.7083
22.9032 | -35.276 | -33.7143
22.9093 | -35.2764 | -33.7206
22.9154 | -35.2768 | -33.7269
22.9215 | -35.2772 | -33.733
22.9276 | -35.2777 | -33.7392
22.9337 | -35.2781 | -33.7449
22.9398 | -35.2785 | -33.7514
22.9459 | -35.279 | -33.7577
22.952 | -35.2794 | -33.7641
22.9581 | -35.2798 | -33.7703
22.9642 | -35.2802 | -33.7764
22.9703 | -35.2807 | -33.7823
22.9764 | -35.2811 | -33.7884
22.9825 | -35.2815 | -33.7943
22.9886 | -35.2819 | -33.8008
22.9948 | -35.2823 | -33.8071
23.0009 | -35.2827 | -33.8131
23.007 | -35.2832 | -33.8197
23.0131 | -35.2837 | -33.8258
23.0192 | -35.284 | -33.8318
23.0253 | -35.2845 | -33.8384
23.0314 | -35.2849 | -33.8446
23.0375 | -35.2853 | -33.8509
23.0436 | -35.2857 | -33.8569
23.0497 | -35.2862 | -33.863
23.0558 | -35.2866 | -33.8694
23.0619 | -35.287 | -33.8756
23.068 | -35.2875 | -33.8819
23.0741 | -35.288 | -33.8873
23.0802 | -35.2884 | -33.8934
23.0863 | -35.2888 | -33.8997
23.0924 | -35.2892 | -33.9059
23.0985 | -35.2896 | -33.912
23.1046 | -35.2901 | -33.9182
23.1107 | -35.2904 | -33.9242
23.1168 | -35.2908 | -33.9305
23.1229 | -35.2913 | -33.9367
23.129 | -35.2917 | -33.9431
23.1351 | -35.2922 | -33.9496
23.1413 | -35.2926 | -33.9556
23.1474 | -35.293 | -33.9615
23.1535 | -35.2935 | -33.9679
23.1596 | -35.2939 | -33.974
23.1657 | -35.2943 | -33.98
23.1718 | -35.2947 | -33.9863
23.1779 | -35.2951 | -33.9927
23.184 | -35.2954 | -33.9989
23.1901 | -35.2959 | -34.0053
23.1962 | -35.2963 | -34.0114
23.2023 | -35.2968 | -34.0174
23.2084 | -35.2972 | -34.0236
23.2145 | -35.2977 | -34.0296
23.2206 | -35.2981 | -34.0358
23.2267 | -35.2985 | -34.0419
23.2328 | -35.2989 | -34.0479
23.2389 | -35.2994 | -34.0544
23.245 | -35.2998 | -34.0609
23.2511 | -35.3002 | -34.067
23.2572 | -35.3007 | -34.0733
23.2633 | -35.3011 | -34.0792
23.2694 | -35.3015 | -34.0852
23.2755 | -35.3019 | -34.0915
23.2817 | -35.3023 | -34.0976
23.2878 | -35.3027 | -34.1035
23.2939 | -35.3032 | -34.1096
23.3 | -35.3036 | -34.1156
23.3061 | -35.304 | -34.1215
23.3122 | -35.3045 | -34.1275
23.3183 | -35.3048 | -34.134
23.3244 | -35.3052 | -34.1398
23.3305 | -35.3056 | -34.1461
23.3366 | -35.306 | -34.1527
23.3427 | -35.3064 | -34.1588
23.3488 | -35.3068 | -34.165
23.3549 | -35.3072 | -34.1713
23.361 | -35.3077 | -34.177
23.3671 | -35.3082 | -34.1831
23.3732 | -35.3087 | -34.189
23.3793 | -35.3091 | -34.1951
23.3854 | -35.3095 | -34.201
23.3915 | -35.3099 | -34.207
23.3976 | -35.3104 | -34.2131
23.4037 | -35.3108 | -34.2192
23.4098 | -35.3112 | -34.2251
23.4159 | -35.3117 | -34.2313
23.422 | -35.3121 | -34.2373
23.4282 | -35.3125 | -34.2431
23.4343 | -35.313 | -34.2493
23.4404 | -35.3133 | -34.2556
23.4465 | -35.3137 | -34.2615
23.4526 | -35.3141 | -34.2679
23.4587 | -35.3145 | -34.2741
23.4648 | -35.3149 | -34.2799
23.4709 | -35.3153 | -34.286
23.477 | -35.3157 | -34.2923
23.4831 | -35.3161 | -34.2982
23.4892 | -35.3166 | -34.3046
23.4953 | -35.317 | -34.3107
23.5014 | -35.3174 | -34.3166
23.5075 | -35.3179 | -34.3227
23.5136 | -35.3183 | -34.329
23.5197 | -35.3187 | -34.3348
23.5258 | -35.3191 | -34.341
23.5319 | -35.3195 | -34.3468
23.538 | -35.32 | -34.3529
23.5441 | -35.3204 | -34.3588
23.5502 | -35.3207 | -34.365
23.5563 | -35.3211 | -34.3711
23.5624 | -35.3215 | -34.377
23.5686 | -35.3219 | -34.3832
23.5747 | -35.3223 | -34.389
23.5808 | -35.3227 | -34.3949
23.5869 | -35.3232 | -34.4008
23.593 | -35.3236 | -34.4068
23.5991 | -35.324 | -34.4129
23.6052 | -35.3245 | -34.4188
23.6113 | -35.3249 | -34.4248
23.6174 | -35.3253 | -34.4308
23.6235 | -35.3258 | -34.4369
23.6296 | -35.3262 | -34.443
23.6357 | -35.3266 | -34.4495
23.6418 | -35.327 | -34.4554
23.6479 | -35.3274 | -34.4614
23.654 | -35.3278 | -34.4673
23.6601 | -35.3283 | -34.4734
23.6662 | -35.3287 | -34.4793
23.6723 | -35.3291 | -34.4852
23.6784 | -35.3295 | -34.491
23.6845 | -35.3299 | -34.4972
23.6906 | -35.3303 | -34.5035
23.6967 | -35.3306 | -34.5097
23.7028 | -35.331 | -34.5159
23.7089 | -35.3314 | -34.5219
23.7151 | -35.3318 | -34.5281
23.7212 | -35.3322 | -34.5339
23.7273 | -35.3326 | -34.54
23.7334 | -35.3331 | -34.5462
23.7395 | -35.3335 | -34.5523
23.7456 | -35.3339 | -34.5587
23.7517 | -35.3342 | -34.5646
23.7578 | -35.3347 | -34.5705
23.7639 | -35.335 | -34.5766
23.77 | -35.3354 | -34.5824
23.7761 | -35.3359 | -34.5888
23.7822 | -35.3362 | -34.595
23.7883 | -35.3365 | -34.6013
23.7944 | -35.3369 | -34.6074
23.8005 | -35.3373 | -34.6133
23.8066 | -35.3377 | -34.6193
23.8127 | -35.3381 | -34.6256
23.8188 | -35.3385 | -34.6318
23.8249 | -35.339 | -34.6379
23.831 | -35.3394 | -34.6443
23.8371 | -35.3398 | -34.65
23.8432 | -35.3402 | -34.6562
23.8493 | -35.3407 | -34.6622
23.8555 | -35.3411 | -34.6684
23.8616 | -35.3415 | -34.6744
23.8677 | -35.3419 | -34.6807
23.8738 | -35.3423 | -34.687
23.8799 | -35.3427 | -34.6929
23.886 | -35.343 | -34.6993
23.8921 | -35.3434 | -34.7056
23.8982 | -35.3438 | -34.7116
23.9043 | -35.3442 | -34.7179
23.9104 | -35.3446 | -34.724
23.9165 | -35.345 | -34.7299
23.9226 | -35.3454 | -34.7357
23.9287 | -35.3458 | -34.7419
23.9348 | -35.3462 | -34.7479
23.9409 | -35.3466 | -34.7541
23.947 | -35.347 | -34.7601
23.9531 | -35.3474 | -34.7656
23.9592 | -35.3478 | -34.7718
23.9653 | -35.3482 | -34.7777
23.9714 | -35.3486 | -34.7837
23.9775 | -35.349 | -34.7897
23.9836 | -35.3494 | -34.7957
23.9897 | -35.3499 | -34.8021
23.9958 | -35.3503 | -34.8084
24.002 | -35.3507 | -34.8146
24.0081 | -35.351 | -34.8206
24.0142 | -35.3514 | -34.827
24.0203 | -35.3518 | -34.8329
24.0264 | -35.3521 | -34.839
24.0325 | -35.3525 | -34.8449
24.0386 | -35.3529 | -34.8511
24.0447 | -35.3534 | -34.8571
24.0508 | -35.3538 | -34.8629
24.0569 | -35.3542 | -34.8694
24.063 | -35.3546 | -34.8753
24.0691 | -35.355 | -34.8812
24.0752 | -35.3554 | -34.8872
24.0813 | -35.3558 | -34.8933
24.0874 | -35.3562 | -34.8994
24.0935 | -35.3566 | -34.9057
24.0996 | -35.357 | -34.9116
24.1057 | -35.3574 | -34.9177
24.1118 | -35.3577 | -34.9239
24.1179 | -35.3581 | -34.9302
24.124 | -35.3585 | -34.936
24.1301 | -35.3589 | -34.9422
24.1362 | -35.3593 | -34.9483
24.1424 | -35.3597 | -34.9542
24.1485 | -35.3601 | -34.9597
24.1546 | -35.3604 | -34.9658
24.1607 | -35.3608 | -34.9716
24.1668 | -35.3612 | -34.9776
24.1729 | -35.3616 | -34.9836
24.179 | -35.362 | -34.9895
24.1851 | -35.3624 | -34.9956
24.1912 | -35.3628 | -35.002
24.1973 | -35.3632 | -35.0082
24.2034 | -35.3636 | -35.0138
24.2095 | -35.364 | -35.0201
24.2156 | -35.3644 | -35.0264
24.2217 | -35.3648 | -35.0327
24.2278 | -35.3653 | -35.0387
24.2339 | -35.3657 | -35.0446
24.24 | -35.366 | -35.0507
24.2461 | -35.3664 | -35.0567
24.2522 | -35.3668 | -35.0627
24.2583 | -35.3672 | -35.0687
24.2644 | -35.3676 | -35.0749
24.2705 | -35.368 | -35.0807
24.2766 | -35.3683 | -35.0869
24.2827 | -35.3687 | -35.0928
24.2889 | -35.3691 | -35.0986
24.295 | -35.3695 | -35.1047
24.3011 | -35.3699 | -35.1107
24.3072 | -35.3703 | -35.1163
24.3133 | -35.3707 | -35.1221
24.3194 | -35.3712 | -35.1285
24.3255 | -35.3716 | -35.1346
24.3316 | -35.372 | -35.1409
24.3377 | -35.3724 | -35.147
24.3438 | -35.3726 | -35.1531
24.3499 | -35.373 | -35.1594
24.356 | -35.3734 | -35.1657
24.3621 | -35.3737 | -35.1719
24.3682 | -35.3741 | -35.1781
24.3743 | -35.3745 | -35.1845
24.3804 | -35.3748 | -35.1898
24.3865 | -35.3752 | -35.1958
24.3926 | -35.3756 | -35.202
24.3987 | -35.376 | -35.2078
24.4048 | -35.3764 | -35.2137
24.4109 | -35.3768 | -35.2202
24.417 | -35.3771 | -35.2258
24.4231 | -35.3775 | -35.232
24.4293 | -35.3779 | -35.2377
24.4354 | -35.3783 | -35.2438
24.4415 | -35.3787 | -35.2497
24.4476 | -35.3791 | -35.2556
24.4537 | -35.3795 | -35.2617
24.4598 | -35.3799 | -35.2675
24.4659 | -35.3803 | -35.2737
24.472 | -35.3806 | -35.2797
24.4781 | -35.381 | -35.2862
24.4842 | -35.3813 | -35.2922
24.4903 | -35.3817 | -35.2982
24.4964 | -35.3821 | -35.304
24.5025 | -35.3825 | -35.3104
24.5086 | -35.3829 | -35.3166
24.5147 | -35.3833 | -35.3223
24.5208 | -35.3836 | -35.3281
24.5269 | -35.384 | -35.3339
24.533 | -35.3844 | -35.3399
24.5391 | -35.3848 | -35.346
24.5452 | -35.3852 | -35.3516
24.5513 | -35.3856 | -35.3576
24.5574 | -35.386 | -35.3638
24.5635 | -35.3863 | -35.3697
24.5696 | -35.3867 | -35.3759
24.5758 | -35.3871 | -35.382
24.5819 | -35.3875 | -35.3879
24.588 | -35.3878 | -35.3936
24.5941 | -35.3882 | -35.3995
24.6002 | -35.3885 | -35.4056
24.6063 | -35.3889 | -35.4118
24.6124 | -35.3893 | -35.418
24.6185 | -35.3896 | -35.424
24.6246 | -35.39 | -35.4302
24.6307 | -35.3905 | -35.4362
24.6368 | -35.3909 | -35.442
24.6429 | -35.3912 | -35.4478
24.649 | -35.3916 | -35.4537
24.6551 | -35.392 | -35.4596
24.6612 | -35.3924 | -35.4653
24.6673 | -35.3928 | -35.4714
24.6734 | -35.3931 | -35.4775
24.6795 | -35.3935 | -35.4835
24.6856 | -35.3939 | -35.4896
24.6917 | -35.3943 | -35.4955
24.6978 | -35.3946 | -35.5014
24.7039 | -35.395 | -35.5075
24.71 | -35.3953 | -35.5135
24.7162 | -35.3957 | -35.5197
24.7223 | -35.396 | -35.5255
24.7284 | -35.3964 | -35.5313
24.7345 | -35.3968 | -35.5374
24.7406 | -35.3972 | -35.5434
24.7467 | -35.3975 | -35.5494
24.7528 | -35.3979 | -35.5553
24.7589 | -35.3983 | -35.561
24.765 | -35.3986 | -35.5667
24.7711 | -35.399 | -35.5729
24.7772 | -35.3994 | -35.5789
24.7833 | -35.3997 | -35.5851
24.7894 | -35.4001 | -35.5909
24.7955 | -35.4004 | -35.5969
24.8016 | -35.4008 | -35.6026
24.8077 | -35.4011 | -35.6088
24.8138 | -35.4016 | -35.6148
24.8199 | -35.4019 | -35.6207
24.826 | -35.4023 | -35.6263
24.8321 | -35.4027 | -35.6324
24.8382 | -35.403 | -35.6385
24.8443 | -35.4034 | -35.6445
24.8504 | -35.4038 | -35.65
24.8565 | -35.4042 | -35.6561
24.8627 | -35.4045 | -35.6621
24.8688 | -35.4049 | -35.6683
24.8749 | -35.4053 | -35.674
24.881 | -35.4057 | -35.6799
24.8871 | -35.4061 | -35.686
24.8932 | -35.4065 | -35.6919
24.8993 | -35.4069 | -35.6977
24.9054 | -35.4073 | -35.7037
24.9115 | -35.4077 | -35.7098
24.9176 | -35.408 | -35.7157
24.9237 | -35.4084 | -35.7217
24.9298 | -35.4087 | -35.7275
24.9359 | -35.4091 | -35.7336
24.942 | -35.4094 | -35.7394
24.9481 | -35.4098 | -35.7454
24.9542 | -35.4101 | -35.7515
24.9603 | -35.4105 | -35.7574
24.9664 | -35.4109 | -35.7634
24.9725 | -35.4113 | -35.7695
24.9786 | -35.4116 | -35.7752
24.9847 | -35.4119 | -35.7809
24.9908 | -35.4123 | -35.7868
24.9969 | -35.4127 | -35.7925
25.0031 | -35.413 | -35.7984
25.0092 | -35.4134 | -35.8042
25.0153 | -35.4138 | -35.8102
25.0214 | -35.4142 | -35.8162
25.0275 | -35.4145 | -35.8221
25.0336 | -35.4149 | -35.8281
25.0397 | -35.4153 | -35.8341
25.0458 | -35.4157 | -35.8399
25.0519 | -35.416 | -35.8453
25.058 | -35.4164 | -35.8513
25.0641 | -35.4168 | -35.857
25.0702 | -35.4171 | -35.8627
25.0763 | -35.4175 | -35.8685
25.0824 | -35.4179 | -35.8742
25.0885 | -35.4182 | -35.88
25.0946 | -35.4186 | -35.8859
25.1007 | -35.419 | -35.8918
25.1068 | -35.4194 | -35.8976
25.1129 | -35.4197 | -35.9036
25.119 | -35.4201 | -35.9095
25.1251 | -35.4204 | -35.9154
25.1312 | -35.4208 | -35.9217
25.1373 | -35.4212 | -35.9276
25.1435 | -35.4215 | -35.9336
25.1496 | -35.4219 | -35.9394
25.1557 | -35.4223 | -35.9452
25.1618 | -35.4226 | -35.9512
25.1679 | -35.423 | -35.9571
25.174 | -35.4233 | -35.9632
25.1801 | -35.4237 | -35.9694
25.1862 | -35.4241 | -35.9752
25.1923 | -35.4244 | -35.9813
25.1984 | -35.4248 | -35.987
25.2045 | -35.4252 | -35.993
25.2106 | -35.4255 | -35.9989
25.2167 | -35.4258 | -36.0049
25.2228 | -35.4261 | -36.0111
25.2289 | -35.4265 | -36.0167
25.235 | -35.4268 | -36.0226
25.2411 | -35.4272 | -36.0287
25.2472 | -35.4276 | -36.0344
25.2533 | -35.4279 | -36.0404
25.2594 | -35.4283 | -36.0465
25.2655 | -35.4286 | -36.0526
25.2716 | -35.429 | -36.0586
25.2777 | -35.4293 | -36.0647
25.2838 | -35.4296 | -36.0708
25.29 | -35.43 | -36.0765
25.2961 | -35.4304 | -36.0823
25.3022 | -35.4307 | -36.0879
25.3083 | -35.4311 | -36.0935
25.3144 | -35.4313 | -36.0993
25.3205 | -35.4317 | -36.1052
25.3266 | -35.4321 | -36.1108
25.3327 | -35.4324 | -36.1165
25.3388 | -35.4327 | -36.1224
25.3449 | -35.4331 | -36.1282
25.351 | -35.4335 | -36.1341
25.3571 | -35.4338 | -36.1399
25.3632 | -35.4341 | -36.1454
25.3693 | -35.4345 | -36.1514
25.3754 | -35.4348 | -36.1574
25.3815 | -35.4352 | -36.163
25.3876 | -35.4356 | -36.169
25.3937 | -35.4359 | -36.1748
25.3998 | -35.4363 | -36.1805
25.4059 | -35.4367 | -36.1863
25.412 | -35.437 | -36.192
25.4181 | -35.4374 | -36.198
25.4242 | -35.4377 | -36.2039
25.4304 | -35.4381 | -36.2098
25.4365 | -35.4384 | -36.2157
25.4426 | -35.4388 | -36.2217
25.4487 | -35.4392 | -36.2278
25.4548 | -35.4395 | -36.2338
25.4609 | -35.4398 | -36.2396
25.467 | -35.4402 | -36.2454
25.4731 | -35.4405 | -36.2513
25.4792 | -35.4409 | -36.2571
25.4853 | -35.4412 | -36.263
25.4914 | -35.4416 | -36.269
25.4975 | -35.442 | -36.2747
25.5036 | -35.4423 | -36.2805
25.5097 | -35.4427 | -36.286
25.5158 | -35.443 | -36.2914
25.5219 | -35.4434 | -36.2972
25.528 | -35.4437 | -36.3029
25.5341 | -35.4441 | -36.3087
25.5402 | -35.4444 | -36.3144
25.5463 | -35.4447 | -36.3206
25.5524 | -35.4451 | -36.3265
25.5585 | -35.4454 | -36.3321
25.5646 | -35.4458 | -36.3379
25.5707 | -35.4461 | -36.3437
25.5769 | -35.4464 | -36.3501
25.583 | -35.4468 | -36.3561
25.5891 | -35.4472 | -36.3622
25.5952 | -35.4475 | -36.3679
25.6013 | -35.4479 | -36.3738
25.6074 | -35.4483 | -36.3795
25.6135 | -35.4486 | -36.385
25.6196 | -35.449 | -36.391
25.6257 | -35.4493 | -36.3966
25.6318 | -35.4496 | -36.4024
25.6379 | -35.4499 | -36.408
25.644 | -35.4503 | -36.4134
25.6501 | -35.4507 | -36.419
25.6562 | -35.451 | -36.4247
25.6623 | -35.4514 | -36.4305
25.6684 | -35.4518 | -36.4363
25.6745 | -35.4521 | -36.4419
25.6806 | -35.4524 | -36.4475
25.6867 | -35.4528 | -36.4534
25.6928 | -35.4531 | -36.4596
25.6989 | -35.4534 | -36.4655
25.705 | -35.4538 | -36.4713
25.7111 | -35.4541 | -36.477
25.7173 | -35.4544 | -36.4827
25.7234 | -35.4548 | -36.4886
25.7295 | -35.4551 | -36.4945
25.7356 | -35.4555 | -36.5003
25.7417 | -35.4558 | -36.506
25.7478 | -35.4562 | -36.5113
25.7539 | -35.4566 | -36.5168
25.76 | -35.4569 | -36.5225
25.7661 | -35.4572 | -36.5282
25.7722 | -35.4576 | -36.5338
25.7783 | -35.4579 | -36.5398
25.7844 | -35.4582 | -36.5452
25.7905 | -35.4585 | -36.5511
25.7966 | -35.4589 | -36.5568
25.8027 | -35.4592 | -36.5626
25.8088 | -35.4595 | -36.5684
25.8149 | -35.4599 | -36.574
25.821 | -35.4602 | -36.5798
25.8271 | -35.4606 | -36.5857
25.8332 | -35.4609 | -36.5916
25.8393 | -35.4613 | -36.5973
25.8454 | -35.4616 | -36.603
25.8515 | -35.462 | -36.6089
25.8576 | -35.4623 | -36.6143
25.8638 | -35.4627 | -36.6203
25.8699 | -35.463 | -36.6262
25.876 | -35.4633 | -36.6321
25.8821 | -35.4636 | -36.6378
25.8882 | -35.464 | -36.6434
25.8943 | -35.4643 | -36.6491
25.9004 | -35.4646 | -36.6549
25.9065 | -35.465 | -36.6606
25.9126 | -35.4653 | -36.6665
25.9187 | -35.4657 | -36.6724
25.9248 | -35.466 | -36.6781
25.9309 | -35.4663 | -36.6841
25.937 | -35.4667 | -36.69
25.9431 | -35.467 | -36.6956
25.9492 | -35.4674 | -36.7011
25.9553 | -35.4677 | -36.7067
25.9614 | -35.4681 | -36.7121
25.9675 | -35.4684 | -36.7175
25.9736 | -35.4688 | -36.7232
25.9797 | -35.4691 | -36.7291
25.9858 | -35.4695 | -36.7348
25.9919 | -35.4698 | -36.7408
25.998 | -35.4701 | -36.7464
26.0042 | -35.4705 | -36.7521
26.0103 | -35.4708 | -36.7575
26.0164 | -35.4712 | -36.7631
26.0225 | -35.4715 | -36.7689
26.0286 | -35.4718 | -36.7748
26.0347 | -35.4721 | -36.7808
26.0408 | -35.4725 | -36.7867
26.0469 | -35.4728 | -36.7922
26.053 | -35.4731 | -36.7981
26.0591 | -35.4735 | -36.8039
26.0652 | -35.4738 | -36.8097
26.0713 | -35.4742 | -36.8155
26.0774 | -35.4745 | -36.8214
26.0835 | -35.4749 | -36.8272
26.0896 | -35.4752 | -36.8331
26.0957 | -35.4755 | -36.8387
26.1018 | -35.4758 | -36.844
26.1079 | -35.4761 | -36.8494
26.114 | -35.4764 | -36.8551
26.1201 | -35.4767 | -36.861
26.1262 | -35.477 | -36.8666
26.1323 | -35.4774 | -36.8719
26.1384 | -35.4777 | -36.8776
26.1445 | -35.4781 | -36.8833
26.1507 | -35.4784 | -36.8891
26.1568 | -35.4787 | -36.8948
26.1629 | -35.479 | -36.9004
26.169 | -35.4794 | -36.906
26.1751 | -35.4796 | -36.9115
26.1812 | -35.48 | -36.9171
26.1873 | -35.4803 | -36.9228
26.1934 | -35.4807 | -36.9285
26.1995 | -35.4811 | -36.9343
26.2056 | -35.4814 | -36.9396
26.2117 | -35.4817 | -36.9454
26.2178 | -35.482 | -36.9509
26.2239 | -35.4824 | -36.9568
26.23 | -35.4827 | -36.9626
26.2361 | -35.483 | -36.9681
26.2422 | -35.4834 | -36.9736
26.2483 | -35.4837 | -36.9791
26.2544 | -35.4841 | -36.9846
26.2605 | -35.4844 | -36.9907
26.2666 | -35.4847 | -36.9959
26.2727 | -35.485 | -37.0018
26.2788 | -35.4853 | -37.0068
26.2849 | -35.4856 | -37.0125
26.2911 | -35.4859 | -37.0177
26.2972 | -35.4863 | -37.0235
26.3033 | -35.4866 | -37.029
26.3094 | -35.4869 | -37.0346
26.3155 | -35.4872 | -37.0401
26.3216 | -35.4876 | -37.0457
26.3277 | -35.4879 | -37.0511
26.3338 | -35.4883 | -37.0566
26.3399 | -35.4886 | -37.0623
26.346 | -35.4889 | -37.0675
26.3521 | -35.4892 | -37.0735
26.3582 | -35.4896 | -37.0789
26.3643 | -35.49 | -37.0846
26.3704 | -35.4903 | -37.0907
26.3765 | -35.4907 | -37.0961
26.3826 | -35.491 | -37.1015
26.3887 | -35.4913 | -37.107
26.3948 | -35.4916 | -37.1126
26.4009 | -35.4919 | -37.1182
26.407 | -35.4922 | -37.1237
26.4131 | -35.4926 | -37.1291
26.4192 | -35.4928 | -37.1347
26.4253 | -35.4932 | -37.1403
26.4314 | -35.4935 | -37.1455
26.4376 | -35.4939 | -37.151
26.4437 | -35.4942 | -37.1567
26.4498 | -35.4945 | -37.1626
26.4559 | -35.4948 | -37.1678
26.462 | -35.4951 | -37.1734
26.4681 | -35.4954 | -37.1791
26.4742 | -35.4958 | -37.1849
26.4803 | -35.496 | -37.1905
26.4864 | -35.4964 | -37.196
26.4925 | -35.4967 | -37.2014
26.4986 | -35.497 | -37.2072
26.5047 | -35.4974 | -37.2129
26.5108 | -35.4977 | -37.218
26.5169 | -35.498 | -37.2236
26.523 | -35.4983 | -37.2292
26.5291 | -35.4986 | -37.2346
26.5352 | -35.4988 | -37.2403
26.5413 | -35.4992 | -37.2462
26.5474 | -35.4996 | -37.2518
26.5535 | -35.4999 | -37.2573
26.5596 | -35.5001 | -37.2628
26.5657 | -35.5005 | -37.2684
26.5718 | -35.5008 | -37.2742
26.578 | -35.5011 | -37.2797
26.5841 | -35.5015 | -37.2856
26.5902 | -35.5018 | -37.2909
26.5963 | -35.5021 | -37.2964
26.6024 | -35.5025 | -37.3022
26.6085 | -35.5028 | -37.308
26.6146 | -35.5031 | -37.3134
26.6207 | -35.5034 | -37.3191
26.6268 | -35.5038 | -37.3248
26.6329 | -35.5041 | -37.3303
26.639 | -35.5044 | -37.3362
26.6451 | -35.5047 | -37.3417
26.6512 | -35.5051 | -37.3471
26.6573 | -35.5054 | -37.3528
26.6634 | -35.5057 | -37.3582
26.6695 | -35.5059 | -37.3637
26.6756 | -35.5062 | -37.3697
26.6817 | -35.5065 | -37.3752
26.6878 | -35.5069 | -37.3808
26.6939 | -35.5072 | -37.3863
26.7 | -35.5074 | -37.3916
26.7061 | -35.5078 | -37.3973
26.7122 | -35.5081 | -37.4029
26.7183 | -35.5084 | -37.4083
26.7245 | -35.5088 | -37.4137
26.7306 | -35.5091 | -37.4189
26.7367 | -35.5094 | -37.4238
26.7428 | -35.5097 | -37.4291
26.7489 | -35.5101 | -37.4349
26.755 | -35.5104 | -37.4405
26.7611 | -35.5107 | -37.4458
26.7672 | -35.511 | -37.4514
26.7733 | -35.5114 | -37.4568
26.7794 | -35.5117 | -37.4625
26.7855 | -35.5121 | -37.468
26.7916 | -35.5124 | -37.4731
26.7977 | -35.5127 | -37.4789
26.8038 | -35.513 | -37.4847
26.8099 | -35.5133 | -37.4903
26.816 | -35.5136 | -37.4959
26.8221 | -35.514 | -37.5012
26.8282 | -35.5143 | -37.5066
26.8343 | -35.5146 | -37.5122
26.8404 | -35.5149 | -37.5174
26.8465 | -35.5153 | -37.5229
26.8526 | -35.5156 | -37.5285
26.8587 | -35.5159 | -37.5339
26.8649 | -35.5162 | -37.5393
26.871 | -35.5165 | -37.5451
26.8771 | -35.5168 | -37.5504
26.8832 | -35.5171 | -37.556
26.8893 | -35.5174 | -37.5614
26.8954 | -35.5178 | -37.5668
26.9015 | -35.5181 | -37.5726
26.9076 | -35.5184 | -37.578
26.9137 | -35.5187 | -37.5829
26.9198 | -35.519 | -37.5885
26.9259 | -35.5193 | -37.5937
26.932 | -35.5196 | -37.5991
26.9381 | -35.5199 | -37.6046
26.9442 | -35.5203 | -37.6097
26.9503 | -35.5205 | -37.6152
26.9564 | -35.5209 | -37.6205
26.9625 | -35.5212 | -37.6261
26.9686 | -35.5214 | -37.6317
26.9747 | -35.5217 | -37.6373
26.9808 | -35.522 | -37.6422
26.9869 | -35.5223 | -37.6476
26.993 | -35.5226 | -37.6534
26.9991 | -35.523 | -37.6584
27.0052 | -35.5233 | -37.6638
27.0114 | -35.5236 | -37.6693
27.0175 | -35.5239 | -37.6746
27.0236 | -35.5243 | -37.6798
27.0297 | -35.5245 | -37.6855
27.0358 | -35.5248 | -37.6913
27.0419 | -35.5252 | -37.6966
27.048 | -35.5255 | -37.7023
27.0541 | -35.5257 | -37.7075
27.0602 | -35.5259 | -37.7126
27.0663 | -35.5262 | -37.7179
27.0724 | -35.5265 | -37.7232
27.0785 | -35.5269 | -37.7288
27.0846 | -35.5272 | -37.7338
27.0907 | -35.5276 | -37.7391
27.0968 | -35.5279 | -37.7448
27.1029 | -35.5282 | -37.75
27.109 | -35.5285 | -37.7551
27.1151 | -35.5288 | -37.7605
27.1212 | -35.5292 | -37.7655
27.1273 | -35.5295 | -37.7714
27.1334 | -35.5298 | -37.7774
27.1395 | -35.5301 | -37.7829
27.1456 | -35.5304 | -37.7882
27.1518 | -35.5308 | -37.7934
27.1579 | -35.531 | -37.7988
27.164 | -35.5314 | -37.8045
27.1701 | -35.5317 | -37.81
27.1762 | -35.5321 | -37.8153
27.1823 | -35.5324 | -37.8208
27.1884 | -35.5327 | -37.8262
27.1945 | -35.533 | -37.8309
27.2006 | -35.5333 | -37.8363
27.2067 | -35.5336 | -37.8417
27.2128 | -35.5339 | -37.8474
27.2189 | -35.5342 | -37.853
27.225 | -35.5346 | -37.8588
27.2311 | -35.5348 | -37.8639
27.2372 | -35.5351 | -37.8692
27.2433 | -35.5354 | -37.8742
27.2494 | -35.5357 | -37.8799
27.2555 | -35.5361 | -37.8852
27.2616 | -35.5364 | -37.8907
27.2677 | -35.5367 | -37.8961
27.2738 | -35.5369 | -37.9016
27.2799 | -35.5372 | -37.907
27.286 | -35.5375 | -37.9123
27.2921 | -35.5378 | -37.9181
27.2983 | -35.5381 | -37.9228
27.3044 | -35.5385 | -37.9285
27.3105 | -35.5388 | -37.934
27.3166 | -35.5391 | -37.9391
27.3227 | -35.5394 | -37.9443
27.3288 | -35.5397 | -37.9498
27.3349 | -35.54 | -37.9545
27.341 | -35.5403 | -37.9598
27.3471 | -35.5405 | -37.9651
27.3532 | -35.5409 | -37.9702
27.3593 | -35.5412 | -37.9755
27.3654 | -35.5415 | -37.981
27.3715 | -35.5418 | -37.9854
27.3776 | -35.5421 | -37.9908
27.3837 | -35.5424 | -37.9961
27.3898 | -35.5427 | -38.0016
27.3959 | -35.543 | -38.0066
27.402 | -35.5433 | -38.0124
27.4081 | -35.5436 | -38.0176
27.4142 | -35.544 | -38.0229
27.4203 | -35.5442 | -38.028
27.4264 | -35.5446 | -38.0334
27.4325 | -35.5449 | -38.0388
27.4387 | -35.5451 | -38.0439
27.4448 | -35.5454 | -38.0491
27.4509 | -35.5456 | -38.0544
27.457 | -35.546 | -38.0598
27.4631 | -35.5463 | -38.0652
27.4692 | -35.5465 | -38.0704
27.4753 | -35.5469 | -38.0755
27.4814 | -35.5472 | -38.0803
27.4875 | -35.5475 | -38.0856
27.4936 | -35.5479 | -38.0911
27.4997 | -35.5482 | -38.0965
27.5058 | -35.5485 | -38.1013
27.5119 | -35.5488 | -38.1062
27.518 | -35.5491 | -38.1115
27.5241 | -35.5494 | -38.1165
27.5302 | -35.5497 | -38.1221
27.5363 | -35.5499 | -38.1274
27.5424 | -35.5502 | -38.1323
27.5485 | -35.5506 | -38.1374
27.5546 | -35.5509 | -38.1423
27.5607 | -35.5511 | -38.1469
27.5668 | -35.5514 | -38.1517
27.5729 | -35.5518 | -38.1567
27.5791 | -35.552 | -38.1615
27.5852 | -35.5524 | -38.1667
27.5913 | -35.5527 | -38.1719
27.5974 | -35.5529 | -38.1773
27.6035 | -35.5532 | -38.1825
27.6096 | -35.5535 | -38.1875
27.6157 | -35.5537 | -38.1926
27.6218 | -35.554 | -38.198
27.6279 | -35.5543 | -38.2038
27.634 | -35.5546 | -38.2087
27.6401 | -35.5548 | -38.2139
27.6462 | -35.5551 | -38.2188
27.6523 | -35.5554 | -38.2244
27.6584 | -35.5557 | -38.2296
27.6645 | -35.556 | -38.2353
27.6706 | -35.5563 | -38.2405
27.6767 | -35.5567 | -38.2456
27.6828 | -35.5569 | -38.2505
27.6889 | -35.5571 | -38.2557
27.695 | -35.5574 | -38.2617
27.7011 | -35.5578 | -38.2665
27.7072 | -35.5581 | -38.2715
27.7133 | -35.5584 | -38.2767
27.7194 | -35.5587 | -38.2807
27.7256 | -35.559 | -38.2857
27.7317 | -35.5593 | -38.2909
27.7378 | -35.5596 | -38.2965
27.7439 | -35.5599 | -38.3021
27.75 | -35.5603 | -38.3073
27.7561 | -35.5606 | -38.3123
27.7622 | -35.5608 | -38.3183
27.7683 | -35.5611 | -38.3232
27.7744 | -35.5616 | -38.3288
27.7805 | -35.5618 | -38.3336
27.7866 | -35.5621 | -38.3388
27.7927 | -35.5623 | -38.3437
27.7988 | -35.5626 | -38.3487
27.8049 | -35.563 | -38.3539
27.811 | -35.5633 | -38.3593
27.8171 | -35.5636 | -38.3647
27.8232 | -35.564 | -38.3702
27.8293 | -35.5641 | -38.3748
27.8354 | -35.5645 | -38.3804
27.8415 | -35.5647 | -38.3854
27.8476 | -35.565 | -38.3907
27.8537 | -35.5652 | -38.396
27.8598 | -35.5655 | -38.4005
27.866 | -35.5658 | -38.4049
27.8721 | -35.5661 | -38.4103
27.8782 | -35.5665 | -38.4155
27.8843 | -35.5669 | -38.4211
27.8904 | -35.5671 | -38.4267
27.8965 | -35.5673 | -38.4319
27.9026 | -35.5675 | -38.4375
27.9087 | -35.5678 | -38.4427
27.9148 | -35.5681 | -38.4476
27.9209 | -35.5684 | -38.4526
27.927 | -35.5686 | -38.4578
27.9331 | -35.5689 | -38.4625
27.9392 | -35.5691 | -38.4684
27.9453 | -35.5694 | -38.4737
27.9514 | -35.5696 | -38.4787
27.9575 | -35.5698 | -38.485
27.9636 | -35.5701 | -38.4896
27.9697 | -35.5705 | -38.4936
27.9758 | -35.571 | -38.4985
27.9819 | -35.5713 | -38.5035
27.988 | -35.5715 | -38.5085
27.9941 | -35.5718 | -38.5139
28.0002 | -35.572 | -38.5184
28.0063 | -35.5722 | -38.5233
28.0125 | -35.5725 | -38.528
28.0186 | -35.5727 | -38.5335
28.0247 | -35.573 | -38.5384
28.0308 | -35.5734 | -38.5432
28.0369 | -35.5735 | -38.5476
28.043 | -35.5737 | -38.553
28.0491 | -35.574 | -38.5582
28.0552 | -35.5742 | -38.5632
28.0613 | -35.5745 | -38.5687
28.0674 | -35.5748 | -38.5731
28.0735 | -35.5751 | -38.5783

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5004 | -42.6325
21.4908 | -35.4875 | -35.3653
16.4948 | -34.8633 | -27.9121
12.6602 | -34.5002 | -21.5778
9.7171 | -34.2739 | -16.6299
7.4582 | -34.1272 | -12.6154
5.7244 | -34.0229 | -9.3579
4.3936 | -33.9562 | -6.7416
3.3722 | -33.9345 | -4.6756
2.5883 | -33.9364 | -3.107
1.9866 | -33.9492 | -1.856
1.5248 | -33.9495 | -0.85441
1.1703 | -33.96 | -0.20066
1 | -33.9761 | 0.12124
0.89822 | -33.9711 | 0.39624
0.68941 | -33.9666 | 0.97654
0.52919 | -33.9549 | 1.4293
0.40612 | -33.9509 | 2.0339
0.31171 | -33.9425 | 2.4835
0.23928 | -33.9509 | 3.2061
0.18358 | -33.9753 | 3.6259
0.14091 | -33.9891 | 4.5047
0.10819 | -33.9977 | 5.6721
0.083017 | -34.0214 | 7.1528
0.063705 | -34.0853 | 9.02
0.048923 | -34.1539 | 11.5176
0.037479 | -34.2785 | 14.7104
0.028801 | -34.4582 | 18.7249
0.022078 | -34.7746 | 24.0635
0.017023 | -35.2718 | 30.458
0.013018 | -36.0478 | 38.1357
0.010014 | -37.1524 | 46.0341

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003689 | -18.7474
0.007377 | -35.1692
0.011066 | -39.5172
0.014754 | -41.9606
0.018443 | -44.2781
0.022131 | -46.3568
0.02582 | -48.0917
0.029508 | -49.5658
0.033197 | -51.6068
0.036885 | -53.0065
0.040574 | -53.6662
0.044262 | -54.1712
0.047951 | -54.7432
0.051639 | -55.9922
0.055328 | -56.9995
0.059016 | -57.6747
0.062705 | -59.3369
0.066393 | -60.4979
0.070082 | -60.4069
0.07377 | -60.3395
0.077459 | -60.6735
0.081147 | -60.4739
0.084836 | -60.7207
0.088524 | -61.3918
0.092213 | -62.438
0.095901 | -63.1783
0.09959 | -62.7379
0.10328 | -63.0409
0.10697 | -64.3956
0.11066 | -64.2582
0.11434 | -65.0234
0.11803 | -64.6236
0.12172 | -64.993
0.12541 | -66.1774
0.1291 | -67.3405
0.13279 | -67.3226
0.13648 | -66.9266
0.14016 | -65.9213
0.14385 | -66.695
0.14754 | -68.6677
0.15123 | -69.1103
0.15492 | -68.2234
0.15861 | -67.8925
0.16229 | -68.0964
0.16598 | -68.3079
0.16967 | -69.4476
0.17336 | -68.2156
0.17705 | -68.8925
0.18074 | -70.0517
0.18443 | -70.1475
0.18811 | -70.3399
0.1918 | -70.3036
0.19549 | -70.7833
0.19918 | -70.0401
0.20287 | -69.9538
0.20656 | -71.0411
0.21024 | -71.8292
0.21393 | -71.5328
0.21762 | -71.9577
0.22131 | -71.0544
0.225 | -70.8128
0.22869 | -71.7669
0.23238 | -72.7009
0.23606 | -74.1779
0.23975 | -73.4883
0.24344 | -72.9664
0.24713 | -73.3401
0.25082 | -72.987
0.25451 | -72.8397
0.2582 | -73.5736
0.26188 | -74.481
0.26557 | -74.2478
0.26926 | -74.6642
0.27295 | -74.4044
0.27664 | -74.3171
0.28033 | -74.4814
0.28402 | -74.4694
0.2877 | -75.3038
0.29139 | -74.3626
0.29508 | -74.4613
0.29877 | -74.7096
0.30246 | -76.4583
0.30615 | -76.5151
0.30983 | -75.7578
0.31352 | -75.9664
0.31721 | -75.7218
0.3209 | -75.5879
0.32459 | -76.3005
0.32828 | -76.496
0.33197 | -75.8431
0.33565 | -76.0309
0.33934 | -75.705
0.34303 | -76.3035
0.34672 | -76.1419
0.35041 | -76.1514
0.3541 | -77.0537
0.35779 | -76.3668
0.36147 | -76.312
0.36516 | -78.0526
0.36885 | -78.5098
0.37254 | -78.1482
0.37623 | -77.2862
0.37992 | -76.8783
0.3836 | -76.972
0.38729 | -77.6152
0.39098 | -77.818
0.39467 | -77.8662
0.39836 | -78.5396
0.40205 | -78.1103
0.40574 | -78.3404
0.40942 | -78.4625
0.41311 | -78.9982
0.4168 | -79.2759
0.42049 | -77.9211
0.42418 | -77.5343
0.42787 | -78.3534
0.43156 | -78.9425
0.43524 | -78.5714
0.43893 | -77.762
0.44262 | -78.8683
0.44631 | -78.5306
0.45 | -78.4077
0.45369 | -79.2997
0.45737 | -80.0758
0.46106 | -80.3788
0.46475 | -79.4092
0.46844 | -78.7886
0.47213 | -80.2072
0.47582 | -80.3288
0.47951 | -79.6393
0.48319 | -79.2228
0.48688 | -79.0515
0.49057 | -79.3361
0.49426 | -79.3027
0.49795 | -79.8726
0.50164 | -81.4304
0.50533 | -81.3697
0.50901 | -81.1087
0.5127 | -80.0681
0.51639 | -79.7093
0.52008 | -80.2677
0.52377 | -80.1528
0.52746 | -80.3
0.53114 | -80.3441
0.53483 | -80.8164
0.53852 | -80.9962
0.54221 | -80.812
0.5459 | -80.6879
0.54959 | -80.6024
0.55328 | -80.562
0.55696 | -80.7944
0.56065 | -81.2639
0.56434 | -81.6587
0.56803 | -81.5537
0.57172 | -81.045
0.57541 | -80.7949
0.5791 | -80.8249
0.58278 | -80.8992
0.58647 | -81.1593
0.59016 | -81.8896
0.59385 | -82.3421
0.59754 | -81.9944
0.60123 | -81.9328
0.60491 | -82.3117
0.6086 | -82.4478
0.61229 | -82.2908
0.61598 | -82.3222
0.61967 | -82.5017
0.62336 | -82.2676
0.62705 | -82.1051
0.63073 | -82.2805
0.63442 | -82.4682
0.63811 | -82.4221
0.6418 | -82.6907
0.64549 | -83.0402
0.64918 | -83.169
0.65287 | -82.8806
0.65655 | -82.4587
0.66024 | -82.1483
0.66393 | -82.0638
0.66762 | -82.3786
0.67131 | -83.0793
0.675 | -83.8278
0.67868 | -83.7256
0.68237 | -83.2344
0.68606 | -82.9641
0.68975 | -82.5005
0.69344 | -82.3159
0.69713 | -82.8686
0.70082 | -83.7611
0.70451 | -83.6099
0.70819 | -82.6939
0.71188 | -82.2263
0.71557 | -82.2903
0.71926 | -82.6196
0.72295 | -82.8238
0.72664 | -82.9458
0.73032 | -83.4226
0.73401 | -84.0639
0.7377 | -84.5403
0.74139 | -84.5281
0.74508 | -83.6747
0.74877 | -82.8243
0.75245 | -82.5609
0.75614 | -83.0462
0.75983 | -83.5347
0.76352 | -83.4526
0.76721 | -83.7783
0.7709 | -84.2204
0.77459 | -84.1276
0.77828 | -84.1759
0.78196 | -84.4684
0.78565 | -84.6434
0.78934 | -84.2058
0.79303 | -83.788
0.79672 | -83.888
0.80041 | -83.8969
0.80409 | -83.6856
0.80778 | -83.6881
0.81147 | -84.0459
0.81516 | -84.6966
0.81885 | -85.128
0.82254 | -85.3234
0.82622 | -85.5599
0.82991 | -85.7502
0.8336 | -85.7803
0.83729 | -85.4689
0.84098 | -85.4486
0.84467 | -85.4588
0.84836 | -85.4602
0.85205 | -85.5277
0.85573 | -85.5112
0.85942 | -85.2254
0.86311 | -84.9606
0.8668 | -84.736
0.87049 | -84.2035
0.87418 | -84.0216
0.87786 | -84.3895
0.88155 | -85.1051
0.88524 | -86.0968
0.88893 | -86.4958
0.89262 | -86.4477
0.89631 | -86.3356
0.89999 | -85.9694
0.90368 | -85.6013
0.90737 | -85.5893
0.91106 | -85.861
0.91475 | -86.1727
0.91844 | -86.5321
0.92213 | -86.9654
0.92582 | -87.0334
0.9295 | -86.9874
0.93319 | -86.9391
0.93688 | -86.739
0.94057 | -86.2189
0.94426 | -86.2444
0.94795 | -86.3134
0.95163 | -85.9164
0.95532 | -85.7778
0.95901 | -86.0327
0.9627 | -86.4909
0.96639 | -86.9149
0.97008 | -87.1356
0.97377 | -86.8636
0.97745 | -86.4598
0.98114 | -86.4309
0.98483 | -86.4566
0.98852 | -86.5667
0.99221 | -87.1266
0.9959 | -87.6332
0.99958 | -88.0404
1.0033 | -88.1373
1.007 | -87.5055
1.0107 | -86.7331
1.0143 | -86.3094
1.018 | -86.4472
1.0217 | -87.1819
1.0254 | -87.6178
1.0291 | -87.5802
1.0328 | -87.5402
1.0365 | -87.3659
1.0402 | -87.105
1.0438 | -86.9131
1.0475 | -87.1119
1.0512 | -87.1922
1.0549 | -86.9646
1.0586 | -87.1709
1.0623 | -87.3474
1.066 | -87.4687
1.0697 | -87.5846
1.0734 | -87.6493
1.077 | -87.6279
1.0807 | -87.6102
1.0844 | -87.6341
1.0881 | -87.6703
1.0918 | -87.7058
1.0955 | -87.7207
1.0992 | -87.6713
1.1029 | -87.514
1.1066 | -87.2961
1.1102 | -87.1397
1.1139 | -87.0751
1.1176 | -87.0123
1.1213 | -86.9758
1.125 | -87.1553
1.1287 | -87.5228
1.1324 | -87.8084
1.1361 | -87.9723
1.1397 | -88.1536
1.1434 | -88.309
1.1471 | -88.281
1.1508 | -88.1344
1.1545 | -88.0107
1.1582 | -87.864
1.1619 | -87.7103
1.1656 | -87.6588
1.1693 | -87.7374
1.1729 | -87.9221
1.1766 | -88.1055
1.1803 | -88.3248
1.184 | -88.7464
1.1877 | -89.0888
1.1914 | -88.8976
1.1951 | -88.4483
1.1988 | -88.1775
1.2025 | -88.1577
1.2061 | -88.3126
1.2098 | -88.4982
1.2135 | -88.542
1.2172 | -88.5192
1.2209 | -88.6454
1.2246 | -88.8534
1.2283 | -88.8441
1.232 | -88.5842
1.2356 | -88.3839
1.2393 | -88.5414
1.243 | -89.0353
1.2467 | -89.4814
1.2504 | -89.5451
1.2541 | -89.2915
1.2578 | -88.9233
1.2615 | -88.612
1.2652 | -88.5899
1.2688 | -88.903
1.2725 | -89.1535
1.2762 | -89.0448
1.2799 | -88.8393
1.2836 | -88.7231
1.2873 | -88.6719
1.291 | -88.6655
1.2947 | -88.6326
1.2984 | -88.6394
1.302 | -88.8693
1.3057 | -89.2342
1.3094 | -89.4506
1.3131 | -89.4874
1.3168 | -89.4598
1.3205 | -89.4228
1.3242 | -89.4614
1.3279 | -89.6056
1.3316 | -89.8238
1.3352 | -90.0541
1.3389 | -90.113
1.3426 | -89.8927
1.3463 | -89.6465
1.35 | -89.6523
1.3537 | -89.8011
1.3574 | -89.7298
1.3611 | -89.4444
1.3647 | -89.2641
1.3684 | -89.3606
1.3721 | -89.569
1.3758 | -89.6207
1.3795 | -89.5947
1.3832 | -89.7017
1.3869 | -89.7724
1.3906 | -89.5905
1.3943 | -89.3739
1.3979 | -89.4171
1.4016 | -89.5902
1.4053 | -89.6604
1.409 | -89.7185
1.4127 | -89.909
1.4164 | -90.1868
1.4201 | -90.3652
1.4238 | -90.3442
1.4275 | -90.239
1.4311 | -90.1011
1.4348 | -89.9333
1.4385 | -89.822
1.4422 | -89.7826
1.4459 | -89.7107
1.4496 | -89.6269
1.4533 | -89.6961
1.457 | -90.0576
1.4606 | -90.6712
1.4643 | -91.1071
1.468 | -90.8924
1.4717 | -90.3214
1.4754 | -89.9549
1.4791 | -89.9163
1.4828 | -89.9589
1.4865 | -89.9343
1.4902 | -89.9829
1.4938 | -90.2775
1.4975 | -90.8045
1.5012 | -91.3171
1.5049 | -91.5073
1.5086 | -91.3315
1.5123 | -91.067
1.516 | -90.963
1.5197 | -90.9043
1.5234 | -90.672
1.527 | -90.3573
1.5307 | -90.0672
1.5344 | -89.8663
1.5381 | -89.9058
1.5418 | -90.2287
1.5455 | -90.6788
1.5492 | -90.8583
1.5529 | -90.521
1.5565 | -90.1211
1.5602 | -90.145
1.5639 | -90.5384
1.5676 | -90.8101
1.5713 | -90.5635
1.575 | -90.038
1.5787 | -89.791
1.5824 | -90.045
1.5861 | -90.4668
1.5897 | -90.5958
1.5934 | -90.4057
1.5971 | -90.3379
1.6008 | -90.3209
1.6045 | -90.3682
1.6082 | -90.4428
1.6119 | -90.5254
1.6156 | -90.6572
1.6193 | -90.8404
1.6229 | -91.0219
1.6266 | -91.1344
1.6303 | -91.1995
1.634 | -91.2778
1.6377 | -91.3153
1.6414 | -91.2048
1.6451 | -90.9358
1.6488 | -90.6659
1.6525 | -90.5928
1.6561 | -90.7839
1.6598 | -91.154
1.6635 | -91.5473
1.6672 | -91.8317
1.6709 | -91.942
1.6746 | -91.8747
1.6783 | -91.7449
1.682 | -91.6993
1.6856 | -91.8139
1.6893 | -92.0333
1.693 | -92.2404
1.6967 | -92.3923
1.7004 | -92.5178
1.7041 | -92.6138
1.7078 | -92.6464
1.7115 | -92.5357
1.7152 | -92.2571
1.7188 | -91.8549
1.7225 | -91.4346
1.7262 | -91.1129
1.7299 | -90.9331
1.7336 | -90.8782
1.7373 | -90.8918
1.741 | -90.9772
1.7447 | -91.2089
1.7484 | -91.5906
1.752 | -92.0266
1.7557 | -92.3624
1.7594 | -92.525
1.7631 | -92.5852
1.7668 | -92.6438
1.7705 | -92.7276
1.7742 | -92.7695
1.7779 | -92.7215
1.7815 | -92.5699
1.7852 | -92.3377
1.7889 | -92.1194
1.7926 | -91.9575
1.7963 | -91.8185
1.8 | -91.6894
1.8037 | -91.615
1.8074 | -91.6482
1.8111 | -91.7855
1.8147 | -91.9942
1.8184 | -92.191
1.8221 | -92.2881
1.8258 | -92.2912
1.8295 | -92.2574
1.8332 | -92.2005
1.8369 | -92.136
1.8406 | -92.1303
1.8443 | -92.2425
1.8479 | -92.4607
1.8516 | -92.6943
1.8553 | -92.835
1.859 | -92.8481
1.8627 | -92.8073
1.8664 | -92.7631
1.8701 | -92.656
1.8738 | -92.4342
1.8774 | -92.1787
1.8811 | -91.9729
1.8848 | -91.8229
1.8885 | -91.7416
1.8922 | -91.7736
1.8959 | -91.9124
1.8996 | -92.0823
1.9033 | -92.2759
1.907 | -92.5636
1.9106 | -92.9797
1.9143 | -93.4016
1.918 | -93.5989
1.9217 | -93.449
1.9254 | -93.0911
1.9291 | -92.747
1.9328 | -92.541
1.9365 | -92.4891
1.9402 | -92.5556
1.9438 | -92.6653
1.9475 | -92.7381
1.9512 | -92.7604
1.9549 | -92.7293
1.9586 | -92.6526
1.9623 | -92.5818
1.966 | -92.5792
1.9697 | -92.6467
1.9734 | -92.7484
1.977 | -92.852
1.9807 | -92.9608
1.9844 | -93.0714
1.9881 | -93.1575
1.9918 | -93.1119
1.9955 | -92.873
1.9992 | -92.5927
2.0029 | -92.4458
2.0065 | -92.4485
2.0102 | -92.543
2.0139 | -92.6943
2.0176 | -92.8711
2.0213 | -93.0094
2.025 | -93.0324
2.0287 | -92.9429
2.0324 | -92.7732
2.0361 | -92.6185
2.0397 | -92.5801
2.0434 | -92.6875
2.0471 | -92.913
2.0508 | -93.1642
2.0545 | -93.3551
2.0582 | -93.4854
2.0619 | -93.5477
2.0656 | -93.5684
2.0693 | -93.5602
2.0729 | -93.5706
2.0766 | -93.6384
2.0803 | -93.7156
2.084 | -93.7132
2.0877 | -93.5619
2.0914 | -93.3419
2.0951 | -93.1768
2.0988 | -93.0503
2.1024 | -92.8869
2.1061 | -92.6962
2.1098 | -92.6018
2.1135 | -92.6381
2.1172 | -92.705
2.1209 | -92.6871
2.1246 | -92.5963
2.1283 | -92.6477
2.132 | -92.7297
2.1356 | -92.8691
2.1393 | -93.0091
2.143 | -93.1064
2.1467 | -93.148
2.1504 | -93.1583
2.1541 | -93.1601
2.1578 | -93.1762
2.1615 | -93.2101
2.1652 | -93.2516
2.1688 | -93.2929
2.1725 | -93.3414
2.1762 | -93.393
2.1799 | -93.4543
2.1836 | -93.5092
2.1873 | -93.5451
2.191 | -93.5756
2.1947 | -93.6202
2.1983 | -93.6877
2.202 | -93.736
2.2057 | -93.7325
2.2094 | -93.6824
2.2131 | -93.6294
2.2168 | -93.6334
2.2205 | -93.7044
2.2242 | -93.8238
2.2279 | -93.9777
2.2315 | -94.1722
2.2352 | -94.3975
2.2389 | -94.606
2.2426 | -94.7639
2.2463 | -94.8346
2.25 | -94.8191
2.2537 | -94.743
2.2574 | -94.6458
2.2611 | -94.5325
2.2647 | -94.3558
2.2684 | -94.0963
2.2721 | -93.7948
2.2758 | -93.5581
2.2795 | -93.4763
2.2832 | -93.5576
2.2869 | -93.7504
2.2906 | -93.9627
2.2943 | -94.123
2.2979 | -94.2031
2.3016 | -94.2127
2.3053 | -94.1539
2.309 | -94.022
2.3127 | -93.8625
2.3164 | -93.7379
2.3201 | -93.6884
2.3238 | -93.7045
2.3274 | -93.7439
2.3311 | -93.7629
2.3348 | -93.7417
2.3385 | -93.6954
2.3422 | -93.6532
2.3459 | -93.6447
2.3496 | -93.6733
2.3533 | -93.739
2.357 | -93.8336
2.3606 | -93.9298
2.3643 | -93.9779
2.368 | -93.9432
2.3717 | -93.8492
2.3754 | -93.7688
2.3791 | -93.7694
2.3828 | -93.8854
2.3865 | -94.071
2.3902 | -94.2249
2.3938 | -94.2537
2.3975 | -94.1777
2.4012 | -94.0692
2.4049 | -93.998
2.4086 | -93.992
2.4123 | -94.0828
2.416 | -94.2807
2.4197 | -94.534
2.4233 | -94.7499
2.427 | -94.8633
2.4307 | -94.8735
2.4344 | -94.7862
2.4381 | -94.6398
2.4418 | -94.4989
2.4455 | -94.4058
2.4492 | -94.3799
2.4529 | -94.4062
2.4565 | -94.4659
2.4602 | -94.5495
2.4639 | -94.6254
2.4676 | -94.6762
2.4713 | -94.6775
2.475 | -94.6199
2.4787 | -94.5014
2.4824 | -94.3915
2.4861 | -94.3574
2.4897 | -94.4304
2.4934 | -94.5692
2.4971 | -94.7171
2.5008 | -94.8088
2.5045 | -94.8434
2.5082 | -94.8759
2.5119 | -94.9602
2.5156 | -95.1041
2.5192 | -95.2836
2.5229 | -95.4479
2.5266 | -95.5438
2.5303 | -95.5165
2.534 | -95.3613
2.5377 | -95.1251
2.5414 | -94.9081
2.5451 | -94.7591
2.5488 | -94.675
2.5524 | -94.638
2.5561 | -94.6228
2.5598 | -94.5955
2.5635 | -94.5365
2.5672 | -94.4524
2.5709 | -94.3488
2.5746 | -94.281
2.5783 | -94.3132
2.582 | -94.4832
2.5856 | -94.7757
2.5893 | -95.1107
2.593 | -95.3696
2.5967 | -95.4811
2.6004 | -95.4532
2.6041 | -95.3472
2.6078 | -95.2525
2.6115 | -95.2196
2.6152 | -95.2573
2.6188 | -95.3778
2.6225 | -95.5918
2.6262 | -95.8832
2.6299 | -96.1864
2.6336 | -96.3986
2.6373 | -96.4335
2.641 | -96.296
2.6447 | -96.0428
2.6483 | -95.7491
2.652 | -95.5056
2.6557 | -95.3613
2.6594 | -95.3731
2.6631 | -95.3478
2.6668 | -95.3383
2.6705 | -95.3157
2.6742 | -95.2729
2.6779 | -95.2356
2.6815 | -95.2152
2.6852 | -95.2162
2.6889 | -95.2411
2.6926 | -95.2836
2.6963 | -95.3348
2.7 | -95.3852
2.7037 | -95.4257
2.7074 | -95.4479
2.7111 | -95.4444
2.7147 | -95.4349
2.7184 | -95.438
2.7221 | -95.4623
2.7258 | -95.4747
2.7295 | -95.471
2.7332 | -95.4352
2.7369 | -95.3757
2.7406 | -95.3246
2.7442 | -95.2935
2.7479 | -95.2999
2.7516 | -95.3416
2.7553 | -95.3972
2.759 | -95.4345
2.7627 | -95.4268
2.7664 | -95.358
2.7701 | -95.2524
2.7738 | -95.1532
2.7774 | -95.0775
2.7811 | -95.0542
2.7848 | -95.0723
2.7885 | -95.1185
2.7922 | -95.1847
2.7959 | -95.2472
2.7996 | -95.3124
2.8033 | -95.3944
2.807 | -95.499
2.8106 | -95.6313
2.8143 | -95.8018
2.818 | -95.9948
2.8217 | -96.1945
2.8254 | -96.3946
2.8291 | -96.55
2.8328 | -96.6223
2.8365 | -96.5901
2.8401 | -96.4596
2.8438 | -96.2545
2.8475 | -96.005
2.8512 | -95.748
2.8549 | -95.4972
2.8586 | -95.2571
2.8623 | -95.0523
2.866 | -94.926
2.8697 | -94.9045
2.8733 | -94.9749
2.877 | -95.1274
2.8807 | -95.3391
2.8844 | -95.5829
2.8881 | -95.8301
2.8918 | -96.0534
2.8955 | -96.2142
2.8992 | -96.2949
2.9029 | -96.331
2.9065 | -96.3454
2.9102 | -96.3539
2.9139 | -96.3522
2.9176 | -96.3359
2.9213 | -96.3315
2.925 | -96.3473
2.9287 | -96.3732
2.9324 | -96.3988
2.9361 | -96.4049
2.9397 | -96.3678
2.9434 | -96.2968
2.9471 | -96.2203
2.9508 | -96.1704
2.9545 | -96.1833
2.9582 | -96.2571
2.9619 | -96.3747
2.9656 | -96.5096
2.9692 | -96.6116
2.9729 | -96.6725
2.9766 | -96.6882
2.9803 | -96.6427
2.984 | -96.5584
2.9877 | -96.4538
2.9914 | -96.3586
2.9951 | -96.2787
2.9988 | -96.2094
3.0024 | -96.1579
3.0061 | -96.1332
3.0098 | -96.1585
3.0135 | -96.2489
3.0172 | -96.4047
3.0209 | -96.595
3.0246 | -96.7921
3.0283 | -96.9541
3.032 | -97.0389
3.0356 | -97.0323
3.0393 | -96.9256
3.043 | -96.7427
3.0467 | -96.5338
3.0504 | -96.344
3.0541 | -96.2197
3.0578 | -96.1854
3.0615 | -96.2311
3.0651 | -96.3184
3.0688 | -96.4272
3.0725 | -96.5407
3.0762 | -96.6512
3.0799 | -96.7401
3.0836 | -96.7973
3.0873 | -96.8045
3.091 | -96.7534
3.0947 | -96.6281
3.0983 | -96.4624
3.102 | -96.292
3.1057 | -96.1529
3.1094 | -96.077
3.1131 | -96.0682
3.1168 | -96.1041
3.1205 | -96.1628
3.1242 | -96.2337
3.1279 | -96.3227
3.1315 | -96.4161
3.1352 | -96.505
3.1389 | -96.588
3.1426 | -96.6492
3.1463 | -96.6498
3.15 | -96.568
3.1537 | -96.3941
3.1574 | -96.1508
3.161 | -95.8908
3.1647 | -95.6741
3.1684 | -95.5649
3.1721 | -95.5822
3.1758 | -95.7067
3.1795 | -95.9061
3.1832 | -96.106
3.1869 | -96.2351
3.1906 | -96.2109
3.1942 | -96.2281
3.1979 | -96.2226
3.2016 | -96.2078
3.2053 | -96.1954
3.209 | -96.197
3.2127 | -96.2215
3.2164 | -96.2578
3.2201 | -96.2955
3.2238 | -96.3359
3.2274 | -96.3877
3.2311 | -96.4601
3.2348 | -96.5751
3.2385 | -96.7216
3.2422 | -96.8737
3.2459 | -97.0139
3.2496 | -97.1115
3.2533 | -97.1475
3.257 | -97.1296
3.2606 | -97.0717
3.2643 | -97.0084
3.268 | -96.9763
3.2717 | -96.9824
3.2754 | -97.0142
3.2791 | -97.0659
3.2828 | -97.1181
3.2865 | -97.1544
3.2901 | -97.1715
3.2938 | -97.173
3.2975 | -97.1636
3.3012 | -97.1381
3.3049 | -97.0967
3.3086 | -97.0574
3.3123 | -97.0228
3.316 | -96.994
3.3197 | -96.9761
3.3233 | -96.963
3.327 | -96.9391
3.3307 | -96.9023
3.3344 | -96.8474
3.3381 | -96.7919
3.3418 | -96.7665
3.3455 | -96.7878
3.3492 | -96.8668
3.3529 | -97.0011
3.3565 | -97.1748
3.3602 | -97.3564
3.3639 | -97.4933
3.3676 | -97.5696
3.3713 | -97.5794
3.375 | -97.5319
3.3787 | -97.4594
3.3824 | -97.3961
3.386 | -97.3486
3.3897 | -97.3264
3.3934 | -97.3271
3.3971 | -97.3366
3.4008 | -97.3473
3.4045 | -97.3503
3.4082 | -97.3287
3.4119 | -97.288
3.4156 | -97.2415
3.4192 | -97.2086
3.4229 | -97.1979
3.4266 | -97.2173
3.4303 | -97.2547
3.434 | -97.2998
3.4377 | -97.3385
3.4414 | -97.3468
3.4451 | -97.3249
3.4488 | -97.2696
3.4524 | -97.2011
3.4561 | -97.1438
3.4598 | -97.1116
3.4635 | -97.1324
3.4672 | -97.2023
3.4709 | -97.2976
3.4746 | -97.4044
3.4783 | -97.4955
3.4819 | -97.5593
3.4856 | -97.5981
3.4893 | -97.6101
3.493 | -97.6021
3.4967 | -97.5985
3.5004 | -97.5962
3.5041 | -97.6118
3.5078 | -97.6345
3.5115 | -97.6657
3.5151 | -97.7164
3.5188 | -97.7791
3.5225 | -97.8635
3.5262 | -97.964
3.5299 | -98.0465
3.5336 | -98.1041
3.5373 | -98.118
3.541 | -98.0942
3.5447 | -98.0483
3.5483 | -97.9828
3.552 | -97.8967
3.5557 | -97.8101
3.5594 | -97.7247
3.5631 | -97.6465
3.5668 | -97.5772
3.5705 | -97.528
3.5742 | -97.5012
3.5779 | -97.5003
3.5815 | -97.5242
3.5852 | -97.5714
3.5889 | -97.6305
3.5926 | -97.6969
3.5963 | -97.7654
3.6 | -97.8408
3.6037 | -97.9215
3.6074 | -97.9782
3.611 | -97.9932
3.6147 | -97.9575
3.6184 | -97.8849
3.6221 | -97.7978
3.6258 | -97.7157
3.6295 | -97.6619
3.6332 | -97.6336
3.6369 | -97.635
3.6406 | -97.6616
3.6442 | -97.7041
3.6479 | -97.7499
3.6516 | -97.7946
3.6553 | -97.8455
3.659 | -97.9049
3.6627 | -97.9804
3.6664 | -98.0537
3.6701 | -98.1049
3.6738 | -98.1096
3.6774 | -98.0579
3.6811 | -97.9477
3.6848 | -97.7976
3.6885 | -97.6408
3.6922 | -97.4982
3.6959 | -97.4058
3.6996 | -97.3886
3.7033 | -97.453
3.7069 | -97.5892
3.7106 | -97.7731
3.7143 | -97.9796
3.718 | -98.1732
3.7217 | -98.2519
3.7254 | -98.347
3.7291 | -98.3928
3.7328 | -98.3899
3.7365 | -98.3382
3.7401 | -98.2547
3.7438 | -98.153
3.7475 | -98.0417
3.7512 | -97.9431
3.7549 | -97.8504
3.7586 | -97.7836
3.7623 | -97.757
3.766 | -97.765
3.7697 | -97.813
3.7733 | -97.8925
3.777 | -97.9978
3.7807 | -98.1139
3.7844 | -98.2103
3.7881 | -98.2871
3.7918 | -98.3336
3.7955 | -98.3412
3.7992 | -98.3057
3.8028 | -98.2398
3.8065 | -98.1652
3.8102 | -98.1013
3.8139 | -98.0486
3.8176 | -98.0253
3.8213 | -98.0235
3.825 | -98.0302
3.8287 | -98.0443
3.8324 | -98.054
3.836 | -98.0575
3.8397 | -98.0514
3.8434 | -98.0371
3.8471 | -98.0218
3.8508 | -98.0051
3.8545 | -97.9934
3.8582 | -97.9825
3.8619 | -97.9764
3.8656 | -97.9684
3.8692 | -97.9474
3.8729 | -97.9196
3.8766 | -97.8968
3.8803 | -97.8656
3.884 | -97.8297
3.8877 | -97.8038
3.8914 | -97.7883
3.8951 | -97.7913
3.8988 | -97.8112
3.9024 | -97.8472
3.9061 | -97.8967
3.9098 | -97.9395
3.9135 | -97.9669
3.9172 | -97.9762
3.9209 | -97.9831
3.9246 | -97.9903
3.9283 | -97.9983
3.9319 | -98.0242
3.9356 | -98.0647
3.9393 | -98.1022
3.943 | -98.1359
3.9467 | -98.1622
3.9504 | -98.178
3.9541 | -98.193
3.9578 | -98.2062
3.9615 | -98.2321
3.9651 | -98.2562
3.9688 | -98.2761
3.9725 | -98.2971
3.9762 | -98.3073
3.9799 | -98.3025
3.9836 | -98.281
3.9873 | -98.2523
3.991 | -98.2337
3.9947 | -98.2196
3.9983 | -98.2196
4.002 | -98.2454
4.0057 | -98.2899
4.0094 | -98.3366
4.0131 | -98.384
4.0168 | -98.4276
4.0205 | -98.465
4.0242 | -98.5001
4.0278 | -98.5477
4.0315 | -98.5983
4.0352 | -98.6418
4.0389 | -98.6791
4.0426 | -98.706
4.0463 | -98.708
4.05 | -98.678
4.0537 | -98.6159
4.0574 | -98.5442
4.061 | -98.4704
4.0647 | -98.4049
4.0684 | -98.3473
4.0721 | -98.3021
4.0758 | -98.2645
4.0795 | -98.2308
4.0832 | -98.2099
4.0869 | -98.1982
4.0906 | -98.1984
4.0942 | -98.2186
4.0979 | -98.2701
4.1016 | -98.3437
4.1053 | -98.4371
4.109 | -98.5497
4.1127 | -98.6647
4.1164 | -98.7707
4.1201 | -98.8518
4.1237 | -98.9061
4.1274 | -98.9419
4.1311 | -98.9589
4.1348 | -98.956
4.1385 | -98.9388
4.1422 | -98.9158
4.1459 | -98.8902
4.1496 | -98.8706
4.1533 | -98.8704
4.1569 | -98.8922
4.1606 | -98.9233
4.1643 | -98.9598
4.168 | -98.9949
4.1717 | -99.0293
4.1754 | -99.0655
4.1791 | -99.0916
4.1828 | -99.1093
4.1865 | -99.113
4.1901 | -99.0923
4.1938 | -99.0522
4.1975 | -98.9856
4.2012 | -98.9104
4.2049 | -98.8321
4.2086 | -98.7524
4.2123 | -98.6812
4.216 | -98.6236
4.2197 | -98.5955
4.2233 | -98.609
4.227 | -98.6612
4.2307 | -98.7313
4.2344 | -98.802
4.2381 | -98.8681
4.2418 | -98.9263
4.2455 | -98.9688
4.2492 | -99.0061
4.2528 | -99.0469
4.2565 | -99.0707
4.2602 | -99.098
4.2639 | -99.1232
4.2676 | -99.1301
4.2713 | -99.1198
4.275 | -99.1066
4.2787 | -99.0866
4.2824 | -99.069
4.286 | -99.0549
4.2897 | -99.0487
4.2934 | -99.0566
4.2971 | -99.0888
4.3008 | -99.1417
4.3045 | -99.2034
4.3082 | -99.2714
4.3119 | -99.3368
4.3156 | -99.3891
4.3192 | -99.4204
4.3229 | -99.4336
4.3266 | -99.4358
4.3303 | -99.4182
4.334 | -99.398
4.3377 | -99.3822
4.3414 | -99.3735
4.3451 | -99.362
4.3487 | -99.3459
4.3524 | -99.33
4.3561 | -99.316
4.3598 | -99.297
4.3635 | -99.2701
4.3672 | -99.2323
4.3709 | -99.1797
4.3746 | -99.1175
4.3783 | -99.0542
4.3819 | -98.9901
4.3856 | -98.9339
4.3893 | -98.9005
4.393 | -98.8929
4.3967 | -98.9013
4.4004 | -98.9307
4.4041 | -98.9791
4.4078 | -99.0429
4.4115 | -99.1173
4.4151 | -99.2026
4.4188 | -99.2966
4.4225 | -99.3728
4.4262 | -99.4271
4.4299 | -99.4677
4.4336 | -99.4919
4.4373 | -99.4926
4.441 | -99.4672
4.4446 | -99.4278
4.4483 | -99.3815
4.452 | -99.3445
4.4557 | -99.3169
4.4594 | -99.3016
4.4631 | -99.3051
4.4668 | -99.3263
4.4705 | -99.3578
4.4742 | -99.3932
4.4778 | -99.4344
4.4815 | -99.4649
4.4852 | -99.4834
4.4889 | -99.4971
4.4926 | -99.5082
4.4963 | -99.5179
4.5 | -99.5254
4.5037 | -99.5348
4.5074 | -99.5559
4.511 | -99.5934
4.5147 | -99.6347
4.5184 | -99.6761
4.5221 | -99.7166
4.5258 | -99.7489
4.5295 | -99.7672
4.5332 | -99.7745
4.5369 | -99.7727
4.5406 | -99.7527
4.5442 | -99.7157
4.5479 | -99.6637
4.5516 | -99.6141
4.5553 | -99.5575
4.559 | -99.5093
4.5627 | -99.4719
4.5664 | -99.4416
4.5701 | -99.4155
4.5737 | -99.3877
4.5774 | -99.3635
4.5811 | -99.3443
4.5848 | -99.3286
4.5885 | -99.3205
4.5922 | -99.3243
4.5959 | -99.3456
4.5996 | -99.3838
4.6033 | -99.4389
4.6069 | -99.5068
4.6106 | -99.5884
4.6143 | -99.6733
4.618 | -99.7577
4.6217 | -99.8377
4.6254 | -99.9107
4.6291 | -99.9708
4.6328 | -100.0249
4.6365 | -100.0706
4.6401 | -100.0968
4.6438 | -100.1107
4.6475 | -100.1015
4.6512 | -100.066
4.6549 | -100.0089
4.6586 | -99.9431
4.6623 | -99.8721
4.666 | -99.8085
4.6696 | -99.7672
4.6733 | -99.7493
4.677 | -99.7524
4.6807 | -99.7783
4.6844 | -99.8226
4.6881 | -99.8751
4.6918 | -99.9268
4.6955 | -99.9836
4.6992 | -100.0362
4.7028 | -100.078
4.7065 | -100.098
4.7102 | -100.1052
4.7139 | -100.0929
4.7176 | -100.0654
4.7213 | -100.0306
4.725 | -99.9832
4.7287 | -99.9229
4.7324 | -99.8626
4.736 | -99.8038
4.7397 | -99.7543
4.7434 | -99.7201
4.7471 | -99.7062
4.7508 | -99.7006
4.7545 | -99.6996
4.7582 | -99.7083
4.7619 | -99.7205
4.7655 | -99.7279
4.7692 | -99.7264
4.7729 | -99.7209
4.7766 | -99.7089
4.7803 | -99.6695
4.784 | -99.608
4.7877 | -99.541
4.7914 | -99.4952
4.7951 | -99.4671
4.7987 | -99.4648
4.8024 | -99.4978
4.8061 | -99.5666
4.8098 | -99.6635
4.8135 | -99.7763
4.8172 | -99.8981
4.8209 | -100.0157
4.8246 | -100.1283
4.8283 | -100.2251
4.8319 | -100.287
4.8356 | -100.3111
4.8393 | -100.3097
4.843 | -100.2825
4.8467 | -100.2289
4.8504 | -100.1614
4.8541 | -100.0953
4.8578 | -100.0439
4.8615 | -100.0147
4.8651 | -100.0026
4.8688 | -100.0025
4.8725 | -100.0179
4.8762 | -100.0449
4.8799 | -100.0679
4.8836 | -100.0908
4.8873 | -100.1136
4.891 | -100.1337
4.8946 | -100.1448
4.8983 | -100.1433
4.902 | -100.1308
4.9057 | -100.1203
4.9094 | -100.1148
4.9131 | -100.1212
4.9168 | -100.1321
4.9205 | -100.1564
4.9242 | -100.2009
4.9278 | -100.2593
4.9315 | -100.3239
4.9352 | -100.3847
4.9389 | -100.4397
4.9426 | -100.4911
4.9463 | -100.5347
4.95 | -100.5639
4.9537 | -100.5722
4.9574 | -100.5601
4.961 | -100.5279
4.9647 | -100.4774
4.9684 | -100.4123
4.9721 | -100.3457
4.9758 | -100.2799
4.9795 | -100.2132
4.9832 | -100.1493
4.9869 | -100.1053
4.9905 | -100.0812
4.9942 | -100.0758
4.9979 | -100.0915
5.0016 | -100.1218
5.0053 | -100.1698
5.009 | -100.2327
5.0127 | -100.2862
5.0164 | -100.3271
5.0201 | -100.3473
5.0237 | -100.346
5.0274 | -100.3276
5.0311 | -100.3022
5.0348 | -100.2811
5.0385 | -100.2626
5.0422 | -100.2544
5.0459 | -100.2579
5.0496 | -100.2692
5.0533 | -100.2935
5.0569 | -100.3212
5.0606 | -100.3502
5.0643 | -100.3731
5.068 | -100.3894
5.0717 | -100.3901
5.0754 | -100.3724
5.0791 | -100.3436
5.0828 | -100.2967
5.0864 | -100.2434
5.0901 | -100.1992
5.0938 | -100.1651
5.0975 | -100.1417
5.1012 | -100.1302
5.1049 | -100.1319
5.1086 | -100.1423
5.1123 | -100.1591
5.116 | -100.1789
5.1196 | -100.2059
5.1233 | -100.2369
5.127 | -100.2727
5.1307 | -100.3189
5.1344 | -100.3755
5.1381 | -100.4311
5.1418 | -100.4871
5.1455 | -100.5489
5.1492 | -100.611
5.1528 | -100.6673
5.1565 | -100.7043
5.1602 | -100.7329
5.1639 | -100.7479
5.1676 | -100.7367
5.1713 | -100.7035
5.175 | -100.6663
5.1787 | -100.629
5.1824 | -100.5894
5.186 | -100.5512
5.1897 | -100.5172
5.1934 | -100.4951
5.1971 | -100.487
5.2008 | -100.4887
5.2045 | -100.4934
5.2082 | -100.5075
5.2119 | -100.5354
5.2155 | -100.5723
5.2192 | -100.6074
5.2229 | -100.636
5.2266 | -100.6653
5.2303 | -100.6838
5.234 | -100.6975
5.2377 | -100.7082
5.2414 | -100.7074
5.2451 | -100.7037
5.2487 | -100.6955
5.2524 | -100.6789
5.2561 | -100.6693
5.2598 | -100.6749
5.2635 | -100.6983
5.2672 | -100.7389
5.2709 | -100.7941
5.2746 | -100.857
5.2783 | -100.9239
5.2819 | -100.9837
5.2856 | -101.0228
5.2893 | -101.0413
5.293 | -101.045
5.2967 | -101.041
5.3004 | -101.0323
5.3041 | -101.0332
5.3078 | -101.0442
5.3114 | -101.0634
5.3151 | -101.0866
5.3188 | -101.1302
5.3225 | -101.155
5.3262 | -101.178
5.3299 | -101.1931
5.3336 | -101.1965
5.3373 | -101.186
5.341 | -101.1614
5.3446 | -101.1255
5.3483 | -101.0794
5.352 | -101.0246
5.3557 | -100.961
5.3594 | -100.9035
5.3631 | -100.8583
5.3668 | -100.8171
5.3705 | -100.7846
5.3742 | -100.762
5.3778 | -100.749
5.3815 | -100.7495
5.3852 | -100.7537
5.3889 | -100.7642
5.3926 | -100.7795
5.3963 | -100.7981
5.4 | -100.8206
5.4037 | -100.8516
5.4073 | -100.878
5.411 | -100.9042
5.4147 | -100.922
5.4184 | -100.9312
5.4221 | -100.9337
5.4258 | -100.9328
5.4295 | -100.9231
5.4332 | -100.9039
5.4369 | -100.8873
5.4405 | -100.8745
5.4442 | -100.8624
5.4479 | -100.8588
5.4516 | -100.8671
5.4553 | -100.8917
5.459 | -100.9245
5.4627 | -100.9657
5.4664 | -101.0085
5.4701 | -101.0496
5.4737 | -101.08
5.4774 | -101.0953
5.4811 | -101.0925
5.4848 | -101.0665
5.4885 | -101.0199
5.4922 | -100.9668
5.4959 | -100.907
5.4996 | -100.8478
5.5033 | -100.7886
5.5069 | -100.7305
5.5106 | -100.6782
5.5143 | -100.6384
5.518 | -100.6084
5.5217 | -100.5897
5.5254 | -100.5909
5.5291 | -100.615
5.5328 | -100.6546
5.5364 | -100.7123
5.5401 | -100.7771
5.5438 | -100.8431
5.5475 | -100.9021
5.5512 | -100.9536
5.5549 | -100.9945
5.5586 | -101.0242
5.5623 | -101.0278
5.566 | -101.0159
5.5696 | -100.9962
5.5733 | -100.9685
5.577 | -100.9345
5.5807 | -100.8934
5.5844 | -100.8594
5.5881 | -100.8313
5.5918 | -100.8062
5.5955 | -100.782
5.5992 | -100.7574
5.6028 | -100.7388
5.6065 | -100.7282
5.6102 | -100.7232
5.6139 | -100.7246
5.6176 | -100.7357
5.6213 | -100.7539
5.625 | -100.7775
5.6287 | -100.8087
5.6323 | -100.8385
5.636 | -100.8573
5.6397 | -100.8631
5.6434 | -100.8569
5.6471 | -100.8451
5.6508 | -100.8171
5.6545 | -100.7816
5.6582 | -100.749
5.6619 | -100.7216
5.6655 | -100.704
5.6692 | -100.6939
5.6729 | -100.6999
5.6766 | -100.7306
5.6803 | -100.7735
5.684 | -100.8282
5.6877 | -100.893
5.6914 | -100.9582
5.6951 | -101.0216
5.6987 | -101.0827
5.7024 | -101.1318
5.7061 | -101.165
5.7098 | -101.1953
5.7135 | -101.216
5.7172 | -101.2262
5.7209 | -101.2323
5.7246 | -101.2374
5.7282 | -101.2357
5.7319 | -101.2213
5.7356 | -101.2074
5.7393 | -101.1917
5.743 | -101.1704
5.7467 | -101.1523
5.7504 | -101.1435
5.7541 | -101.1373
5.7578 | -101.1325
5.7614 | -101.114
5.7651 | -101.0948
5.7688 | -101.0743
5.7725 | -101.0474
5.7762 | -101.0156
5.7799 | -100.9795
5.7836 | -100.9438
5.7873 | -100.9182
5.791 | -100.9037
5.7946 | -100.8917
5.7983 | -100.892
5.802 | -100.9078
5.8057 | -100.9354
5.8094 | -100.9752
5.8131 | -101.0304
5.8168 | -101.1049
5.8205 | -101.185
5.8242 | -101.2726
5.8278 | -101.3662
5.8315 | -101.4594
5.8352 | -101.5378
5.8389 | -101.6007
5.8426 | -101.6476
5.8463 | -101.6768
5.85 | -101.633
5.8537 | -101.6204
5.8573 | -101.5919
5.861 | -101.549
5.8647 | -101.502
5.8684 | -101.4611
5.8721 | -101.4164
5.8758 | -101.3718
5.8795 | -101.3229
5.8832 | -101.2794
5.8869 | -101.2492
5.8905 | -101.2293
5.8942 | -101.2144
5.8979 | -101.2027
5.9016 | -101.1931
5.9053 | -101.187
5.909 | -101.1775
5.9127 | -101.1633
5.9164 | -101.1502
5.9201 | -101.1356
5.9237 | -101.1193
5.9274 | -101.1053
5.9311 | -101.0983
5.9348 | -101.0965
5.9385 | -101.1002
5.9422 | -101.1104
5.9459 | -101.1174
5.9496 | -101.1228
5.9532 | -101.1288
5.9569 | -101.1353
5.9606 | -101.135
5.9643 | -101.1313
5.968 | -101.1233
5.9717 | -101.1174
5.9754 | -101.1092
5.9791 | -101.0971
5.9828 | -101.0873
5.9864 | -101.0852
5.9901 | -101.0947
5.9938 | -101.1158
5.9975 | -101.1455
6.0012 | -101.1922
6.0049 | -101.2526
6.0086 | -101.3187
6.0123 | -101.389
6.016 | -101.4625
6.0196 | -101.531
6.0233 | -101.585
6.027 | -101.6297
6.0307 | -101.669
6.0344 | -101.6998
6.0381 | -101.7255
6.0418 | -101.7421
6.0455 | -101.7551
6.0491 | -101.7645
6.0528 | -101.7756
6.0565 | -101.786
6.0602 | -101.7954
6.0639 | -101.8022
6.0676 | -101.8075
6.0713 | -101.811
6.075 | -101.8138
6.0787 | -101.8055
6.0823 | -101.7814
6.086 | -101.7431
6.0897 | -101.692
6.0934 | -101.6284
6.0971 | -101.5567
6.1008 | -101.4773
6.1045 | -101.4005
6.1082 | -101.3341
6.1119 | -101.2798
6.1155 | -101.2435
6.1192 | -101.2254
6.1229 | -101.2243
6.1266 | -101.2412
6.1303 | -101.2745
6.134 | -101.3245
6.1377 | -101.3891
6.1414 | -101.4629
6.1451 | -101.5423
6.1487 | -101.6242
6.1524 | -101.7089
6.1561 | -101.7905
6.1598 | -101.8629
6.1635 | -101.9249
6.1672 | -101.9777
6.1709 | -102.0267
6.1746 | -102.0689
6.1782 | -102.1007
6.1819 | -102.1271
6.1856 | -102.1412
6.1893 | -102.154
6.193 | -102.1745
6.1967 | -102.191
6.2004 | -102.2097
6.2041 | -102.2255
6.2078 | -102.2413
6.2114 | -102.2539
6.2151 | -102.2561
6.2188 | -102.2564
6.2225 | -102.2633
6.2262 | -102.2558
6.2299 | -102.2357
6.2336 | -102.2087
6.2373 | -102.1684
6.241 | -102.1252
6.2446 | -102.0707
6.2483 | -102.0144
6.252 | -101.9605
6.2557 | -101.9131
6.2594 | -101.8698
6.2631 | -101.832
6.2668 | -101.8077
6.2705 | -101.802
6.2741 | -101.7996
6.2778 | -101.8056
6.2815 | -101.8164
6.2852 | -101.8244
6.2889 | -101.8359
6.2926 | -101.8436
6.2963 | -101.8513
6.3 | -101.8551
6.3037 | -101.8608
6.3073 | -101.8706
6.311 | -101.8782
6.3147 | -101.8858
6.3184 | -101.9027
6.3221 | -101.9165
6.3258 | -101.9289
6.3295 | -101.9457
6.3332 | -101.9612
6.3369 | -101.9726
6.3405 | -101.9722
6.3442 | -101.9635
6.3479 | -101.9442
6.3516 | -101.913
6.3553 | -101.8794
6.359 | -101.839
6.3627 | -101.7957
6.3664 | -101.7569
6.37 | -101.7228
6.3737 | -101.69
6.3774 | -101.6663
6.3811 | -101.6389
6.3848 | -101.6327
6.3885 | -101.6226
6.3922 | -101.6198
6.3959 | -101.6249
6.3996 | -101.6369
6.4032 | -101.6536
6.4069 | -101.6782
6.4106 | -101.7093
6.4143 | -101.7521
6.418 | -101.8071
6.4217 | -101.8615
6.4254 | -101.9161
6.4291 | -101.967
6.4328 | -102.014
6.4364 | -102.0626
6.4401 | -102.103
6.4438 | -102.1305
6.4475 | -102.1531
6.4512 | -102.1566
6.4549 | -102.1488
6.4586 | -102.1342
6.4623 | -102.1112
6.466 | -102.0746
6.4696 | -102.0334
6.4733 | -101.9908
6.477 | -101.9616
6.4807 | -101.9319
6.4844 | -101.8998
6.4881 | -101.8755
6.4918 | -101.8557
6.4955 | -101.8373
6.4991 | -101.828
6.5028 | -101.8272
6.5065 | -101.8205
6.5102 | -101.8147
6.5139 | -101.8119
6.5176 | -101.8164
6.5213 | -101.82
6.525 | -101.8233
6.5287 | -101.8285
6.5323 | -101.831
6.536 | -101.8242
6.5397 | -101.8054
6.5434 | -101.7889
6.5471 | -101.7691
6.5508 | -101.7515
6.5545 | -101.7271
6.5582 | -101.7018
6.5619 | -101.6826
6.5655 | -101.669
6.5692 | -101.6615
6.5729 | -101.6599
6.5766 | -101.6654
6.5803 | -101.6832
6.584 | -101.7147
6.5877 | -101.7544
6.5914 | -101.8088
6.595 | -101.8706
6.5987 | -101.9323
6.6024 | -101.9973
6.6061 | -102.0737
6.6098 | -102.15
6.6135 | -102.2207
6.6172 | -102.2786
6.6209 | -102.3355
6.6246 | -102.3817
6.6282 | -102.4107
6.6319 | -102.431
6.6356 | -102.4432
6.6393 | -102.4445
6.643 | -102.4259
6.6467 | -102.3963
6.6504 | -102.3581
6.6541 | -102.3224
6.6578 | -102.2874
6.6614 | -102.2554
6.6651 | -102.2277
6.6688 | -102.2009
6.6725 | -102.1812
6.6762 | -102.1677
6.6799 | -102.1612
6.6836 | -102.1633
6.6873 | -102.1757
6.6909 | -102.191
6.6946 | -102.2149
6.6983 | -102.2462
6.702 | -102.2808
6.7057 | -102.3101
6.7094 | -102.3349
6.7131 | -102.3544
6.7168 | -102.3686
6.7205 | -102.3744
6.7241 | -102.3695
6.7278 | -102.3555
6.7315 | -102.3363
6.7352 | -102.314
6.7389 | -102.2989
6.7426 | -102.2898
6.7463 | -102.2909
6.75 | -102.3049
6.7537 | -102.326
6.7573 | -102.3572
6.761 | -102.3975
6.7647 | -102.4412
6.7684 | -102.4884
6.7721 | -102.5287
6.7758 | -102.5542
6.7795 | -102.5712
6.7832 | -102.5758
6.7869 | -102.5625
6.7905 | -102.5337
6.7942 | -102.4906
6.7979 | -102.4388
6.8016 | -102.3752
6.8053 | -102.3165
6.809 | -102.2679
6.8127 | -102.2315
6.8164 | -102.2068
6.82 | -102.1913
6.8237 | -102.1937
6.8274 | -102.2127
6.8311 | -102.2357
6.8348 | -102.2688
6.8385 | -102.3039
6.8422 | -102.3352
6.8459 | -102.3684
6.8496 | -102.3997
6.8532 | -102.4237
6.8569 | -102.4404
6.8606 | -102.4491
6.8643 | -102.4457
6.868 | -102.4315
6.8717 | -102.4151
6.8754 | -102.3952
6.8791 | -102.3678
6.8828 | -102.3378
6.8864 | -102.313
6.8901 | -102.2964
6.8938 | -102.2869
6.8975 | -102.2875
6.9012 | -102.2991
6.9049 | -102.3135
6.9086 | -102.3329
6.9123 | -102.3557
6.9159 | -102.387
6.9196 | -102.4125
6.9233 | -102.4411
6.927 | -102.466
6.9307 | -102.4893
6.9344 | -102.513
6.9381 | -102.5296
6.9418 | -102.5434
6.9455 | -102.5585
6.9491 | -102.5634
6.9528 | -102.5652
6.9565 | -102.5715
6.9602 | -102.5755
6.9639 | -102.5729
6.9676 | -102.5763
6.9713 | -102.5752
6.975 | -102.562
6.9787 | -102.5452
6.9823 | -102.5263
6.986 | -102.5024
6.9897 | -102.4771
6.9934 | -102.4512
6.9971 | -102.4235
7.0008 | -102.3915
7.0045 | -102.3669
7.0082 | -102.3498
7.0118 | -102.3332
7.0155 | -102.3225
7.0192 | -102.3192
7.0229 | -102.3276
7.0266 | -102.3401
7.0303 | -102.3561
7.034 | -102.3754
7.0377 | -102.3966
7.0414 | -102.4228
7.045 | -102.4499
7.0487 | -102.476
7.0524 | -102.5007
7.0561 | -102.5228
7.0598 | -102.5376
7.0635 | -102.5437
7.0672 | -102.5497
7.0709 | -102.5585
7.0746 | -102.5573
7.0782 | -102.5526
7.0819 | -102.5473
7.0856 | -102.5298
7.0893 | -102.509
7.093 | -102.4901
7.0967 | -102.463
7.1004 | -102.4344
7.1041 | -102.4109
7.1078 | -102.3878
7.1114 | -102.3654
7.1151 | -102.3408
7.1188 | -102.3212
7.1225 | -102.304
7.1262 | -102.2871
7.1299 | -102.2756
7.1336 | -102.2712
7.1373 | -102.27
7.1409 | -102.2718
7.1446 | -102.2825
7.1483 | -102.2991
7.152 | -102.3199
7.1557 | -102.3443
7.1594 | -102.3687
7.1631 | -102.3942
7.1668 | -102.4233
7.1705 | -102.4522
7.1741 | -102.4711
7.1778 | -102.4869
7.1815 | -102.5063
7.1852 | -102.5286
7.1889 | -102.5522
7.1926 | -102.5806
7.1963 | -102.6053
7.2 | -102.6292
7.2037 | -102.6539
7.2073 | -102.6758
7.211 | -102.6933
7.2147 | -102.7119
7.2184 | -102.7178
7.2221 | -102.7268
7.2258 | -102.7278
7.2295 | -102.7247
7.2332 | -102.7188
7.2368 | -102.7037
7.2405 | -102.683
7.2442 | -102.6653
7.2479 | -102.6423
7.2516 | -102.6127
7.2553 | -102.5851
7.259 | -102.5636
7.2627 | -102.5493
7.2664 | -102.54
7.27 | -102.5397
7.2737 | -102.542
7.2774 | -102.554
7.2811 | -102.5758
7.2848 | -102.5991
7.2885 | -102.6233
7.2922 | -102.651
7.2959 | -102.6792
7.2996 | -102.7066
7.3032 | -102.7244
7.3069 | -102.7341
7.3106 | -102.7498
7.3143 | -102.7617
7.318 | -102.7663
7.3217 | -102.7709
7.3254 | -102.7812
7.3291 | -102.7836
7.3327 | -102.7874
7.3364 | -102.7932
7.3401 | -102.7994
7.3438 | -102.8106
7.3475 | -102.8233
7.3512 | -102.8334
7.3549 | -102.8464
7.3586 | -102.8616
7.3623 | -102.8735
7.3659 | -102.8871
7.3696 | -102.9014
7.3733 | -102.9164
7.377 | -102.9303
7.3807 | -102.9434
7.3844 | -102.9501
7.3881 | -102.9545
7.3918 | -102.9618
7.3955 | -102.9607
7.3991 | -102.9548
7.4028 | -102.9475
7.4065 | -102.94
7.4102 | -102.9362
7.4139 | -102.9325
7.4176 | -102.9208
7.4213 | -102.913
7.425 | -102.9103
7.4287 | -102.9039
7.4323 | -102.8969
7.436 | -102.8914
7.4397 | -102.8929
7.4434 | -102.8963
7.4471 | -102.8976
7.4508 | -102.896
7.4545 | -102.8899
7.4582 | -102.8836
7.4618 | -102.8773
7.4655 | -102.8685
7.4692 | -102.8591
7.4729 | -102.8508
7.4766 | -102.8423
7.4803 | -102.8308
7.484 | -102.8202
7.4877 | -102.8116
7.4914 | -102.8025
7.495 | -102.7996
7.4987 | -102.7972
7.5024 | -102.7919
7.5061 | -102.7956
7.5098 | -102.7987
7.5135 | -102.799
7.5172 | -102.7975
7.5209 | -102.7941
7.5246 | -102.7912
7.5282 | -102.7897
7.5319 | -102.7853
7.5356 | -102.7827
7.5393 | -102.7887
7.543 | -102.8041
7.5467 | -102.8291
7.5504 | -102.8579
7.5541 | -102.8956
7.5577 | -102.9446
7.5614 | -103.0049
7.5651 | -103.0733
7.5688 | -103.1486
7.5725 | -103.223
7.5762 | -103.2936
7.5799 | -103.3528
7.5836 | -103.3951
7.5873 | -103.4279
7.5909 | -103.4539
7.5946 | -103.4702
7.5983 | -103.4748
7.602 | -103.4699
7.6057 | -103.4624
7.6094 | -103.4459
7.6131 | -103.4264
7.6168 | -103.4018
7.6205 | -103.3747
7.6241 | -103.3438
7.6278 | -103.3147
7.6315 | -103.2876
7.6352 | -103.2605
7.6389 | -103.2313
7.6426 | -103.2033
7.6463 | -103.1802
7.65 | -103.1593
7.6536 | -103.1365
7.6573 | -103.1127
7.661 | -103.0905
7.6647 | -103.0753
7.6684 | -103.0608
7.6721 | -103.0542
7.6758 | -103.0533
7.6795 | -103.0518
7.6832 | -103.0584
7.6868 | -103.0701
7.6905 | -103.0916
7.6942 | -103.119
7.6979 | -103.1529
7.7016 | -103.1904
7.7053 | -103.2255
7.709 | -103.2642
7.7127 | -103.302
7.7164 | -103.3393
7.72 | -103.3778
7.7237 | -103.409
7.7274 | -103.4362
7.7311 | -103.4634
7.7348 | -103.4819
7.7385 | -103.4957
7.7422 | -103.4994
7.7459 | -103.4881
7.7496 | -103.4732
7.7532 | -103.4543
7.7569 | -103.4295
7.7606 | -103.4005
7.7643 | -103.372
7.768 | -103.3454
7.7717 | -103.3173
7.7754 | -103.2909
7.7791 | -103.2592
7.7827 | -103.2253
7.7864 | -103.1942
7.7901 | -103.1639
7.7938 | -103.1421
7.7975 | -103.1219
7.8012 | -103.1103
7.8049 | -103.103
7.8086 | -103.0989
7.8123 | -103.0916
7.8159 | -103.0863
7.8196 | -103.0877
7.8233 | -103.0902
7.827 | -103.0954
7.8307 | -103.0979
7.8344 | -103.0976
7.8381 | -103.1059
7.8418 | -103.118
7.8455 | -103.1277
7.8491 | -103.1402
7.8528 | -103.1614
7.8565 | -103.1891
7.8602 | -103.2154
7.8639 | -103.2334
7.8676 | -103.2529
7.8713 | -103.2703
7.875 | -103.2803
7.8786 | -103.2847
7.8823 | -103.2782
7.886 | -103.2634
7.8897 | -103.2419
7.8934 | -103.223
7.8971 | -103.2036
7.9008 | -103.1795
7.9045 | -103.1581
7.9082 | -103.1419
7.9118 | -103.1317
7.9155 | -103.1261
7.9192 | -103.1296
7.9229 | -103.1504
7.9266 | -103.181
7.9303 | -103.2082
7.934 | -103.2385
7.9377 | -103.2744
7.9414 | -103.3102
7.945 | -103.3477
7.9487 | -103.3815
7.9524 | -103.4097
7.9561 | -103.4346
7.9598 | -103.4575
7.9635 | -103.4739
7.9672 | -103.4832
7.9709 | -103.4882
7.9745 | -103.4926
7.9782 | -103.4907
7.9819 | -103.4848
7.9856 | -103.4846
7.9893 | -103.4832
7.993 | -103.4864
7.9967 | -103.4894
8.0004 | -103.5016
8.0041 | -103.522
8.0077 | -103.5452
8.0114 | -103.5701
8.0151 | -103.6026
8.0188 | -103.6384
8.0225 | -103.6758
8.0262 | -103.7161
8.0299 | -103.7624
8.0336 | -103.8054
8.0373 | -103.8363
8.0409 | -103.8538
8.0446 | -103.8633
8.0483 | -103.8643
8.052 | -103.8515
8.0557 | -103.8243
8.0594 | -103.788
8.0631 | -103.7412
8.0668 | -103.6885
8.0704 | -103.6284
8.0741 | -103.5627
8.0778 | -103.4975
8.0815 | -103.436
8.0852 | -103.3782
8.0889 | -103.3279
8.0926 | -103.2867
8.0963 | -103.2558
8.1 | -103.231
8.1036 | -103.2079
8.1073 | -103.196
8.111 | -103.1974
8.1147 | -103.2077
8.1184 | -103.2226
8.1221 | -103.2389
8.1258 | -103.2576
8.1295 | -103.2755
8.1332 | -103.2893
8.1368 | -103.3035
8.1405 | -103.3162
8.1442 | -103.3243
8.1479 | -103.3268
8.1516 | -103.3224
8.1553 | -103.3184
8.159 | -103.3155
8.1627 | -103.3093
8.1664 | -103.3046
8.17 | -103.3011
8.1737 | -103.2959
8.1774 | -103.2937
8.1811 | -103.2959
8.1848 | -103.3048
8.1885 | -103.3171
8.1922 | -103.3314
8.1959 | -103.3502
8.1995 | -103.3719
8.2032 | -103.3964
8.2069 | -103.4229
8.2106 | -103.4513
8.2143 | -103.4824
8.218 | -103.5088
8.2217 | -103.528
8.2254 | -103.552
8.2291 | -103.5698
8.2327 | -103.5837
8.2364 | -103.5915
8.2401 | -103.5991
8.2438 | -103.6048
8.2475 | -103.6069
8.2512 | -103.6086
8.2549 | -103.6112
8.2586 | -103.5999
8.2623 | -103.5858
8.2659 | -103.5752
8.2696 | -103.5638
8.2733 | -103.5492
8.277 | -103.5366
8.2807 | -103.5293
8.2844 | -103.527
8.2881 | -103.5182
8.2918 | -103.5062
8.2954 | -103.4975
8.2991 | -103.4908
8.3028 | -103.4798
8.3065 | -103.4681
8.3102 | -103.4555
8.3139 | -103.4452
8.3176 | -103.4315
8.3213 | -103.4146
8.325 | -103.3948
8.3286 | -103.3758
8.3323 | -103.3581
8.336 | -103.3406
8.3397 | -103.334
8.3434 | -103.334
8.3471 | -103.3244
8.3508 | -103.3184
8.3545 | -103.3158
8.3582 | -103.3232
8.3618 | -103.337
8.3655 | -103.3534
8.3692 | -103.374
8.3729 | -103.3957
8.3766 | -103.4217
8.3803 | -103.4564
8.384 | -103.4897
8.3877 | -103.5237
8.3913 | -103.5566
8.395 | -103.5924
8.3987 | -103.6215
8.4024 | -103.6451
8.4061 | -103.6701
8.4098 | -103.6936
8.4135 | -103.7203
8.4172 | -103.7387
8.4209 | -103.7521
8.4245 | -103.763
8.4282 | -103.7658
8.4319 | -103.763
8.4356 | -103.7504
8.4393 | -103.7257
8.443 | -103.696
8.4467 | -103.6697
8.4504 | -103.6399
8.4541 | -103.6088
8.4577 | -103.5765
8.4614 | -103.5386
8.4651 | -103.4978
8.4688 | -103.4719
8.4725 | -103.4462
8.4762 | -103.4197
8.4799 | -103.3964
8.4836 | -103.371
8.4873 | -103.3518
8.4909 | -103.3377
8.4946 | -103.3258
8.4983 | -103.3153
8.502 | -103.3121
8.5057 | -103.3151
8.5094 | -103.3213
8.5131 | -103.3379
8.5168 | -103.3444
8.5204 | -103.3542
8.5241 | -103.3705
8.5278 | -103.3918
8.5315 | -103.4097
8.5352 | -103.4284
8.5389 | -103.4466
8.5426 | -103.4656
8.5463 | -103.4829
8.55 | -103.5007
8.5536 | -103.5194
8.5573 | -103.5403
8.561 | -103.5638
8.5647 | -103.5866
8.5684 | -103.6055
8.5721 | -103.6167
8.5758 | -103.6319
8.5795 | -103.6507
8.5832 | -103.6741
8.5868 | -103.6987
8.5905 | -103.7252
8.5942 | -103.7474
8.5979 | -103.7705
8.6016 | -103.7972
8.6053 | -103.8212
8.609 | -103.8519
8.6127 | -103.8804
8.6163 | -103.906
8.62 | -103.9366
8.6237 | -103.9595
8.6274 | -103.9764
8.6311 | -103.9936
8.6348 | -104.0068
8.6385 | -104.015
8.6422 | -104.0173
8.6459 | -104.0129
8.6495 | -104.002
8.6532 | -103.9842
8.6569 | -103.9598
8.6606 | -103.9339
8.6643 | -103.9063
8.668 | -103.8741
8.6717 | -103.8412
8.6754 | -103.814
8.6791 | -103.7903
8.6827 | -103.7732
8.6864 | -103.7589
8.6901 | -103.7509
8.6938 | -103.7568
8.6975 | -103.7694
8.7012 | -103.7878
8.7049 | -103.809
8.7086 | -103.8352
8.7122 | -103.8707
8.7159 | -103.9138
8.7196 | -103.958
8.7233 | -104.0024
8.727 | -104.0524
8.7307 | -104.1002
8.7344 | -104.1495
8.7381 | -104.1951
8.7418 | -104.2381
8.7454 | -104.2764
8.7491 | -104.3091
8.7528 | -104.3375
8.7565 | -104.354
8.7602 | -104.3598
8.7639 | -104.3616
8.7676 | -104.3599
8.7713 | -104.3467
8.775 | -104.3323
8.7786 | -104.3108
8.7823 | -104.2863
8.786 | -104.2596
8.7897 | -104.2295
8.7934 | -104.1977
8.7971 | -104.164
8.8008 | -104.1347
8.8045 | -104.1034
8.8082 | -104.0723
8.8118 | -104.0433
8.8155 | -104.0169
8.8192 | -103.9914
8.8229 | -103.9708
8.8266 | -103.9546
8.8303 | -103.9435
8.834 | -103.9373
8.8377 | -103.9319
8.8413 | -103.9249
8.845 | -103.925
8.8487 | -103.929
8.8524 | -103.9371
8.8561 | -103.9453
8.8598 | -103.9567
8.8635 | -103.9731
8.8672 | -103.9925
8.8709 | -104.0154
8.8745 | -104.0415
8.8782 | -104.0658
8.8819 | -104.0923
8.8856 | -104.1211
8.8893 | -104.151
8.893 | -104.1792
8.8967 | -104.204
8.9004 | -104.2274
8.9041 | -104.2415
8.9077 | -104.2497
8.9114 | -104.252
8.9151 | -104.2491
8.9188 | -104.2458
8.9225 | -104.2414
8.9262 | -104.2328
8.9299 | -104.2138
8.9336 | -104.1931
8.9372 | -104.1736
8.9409 | -104.1553
8.9446 | -104.1385
8.9483 | -104.1225
8.952 | -104.1046
8.9557 | -104.0912
8.9594 | -104.0835
8.9631 | -104.0764
8.9668 | -104.069
8.9704 | -104.0659
8.9741 | -104.0664
8.9778 | -104.0687
8.9815 | -104.0642
8.9852 | -104.0605
8.9889 | -104.0605
8.9926 | -104.058
8.9963 | -104.0589
9 | -104.0587
9.0036 | -104.0559
9.0073 | -104.0567
9.011 | -104.0577
9.0147 | -104.0547
9.0184 | -104.0483
9.0221 | -104.0407
9.0258 | -104.0358
9.0295 | -104.0347
9.0331 | -104.0358
9.0368 | -104.0407
9.0405 | -104.0492
9.0442 | -104.0692
9.0479 | -104.0824
9.0516 | -104.0973
9.0553 | -104.1161
9.059 | -104.1358
9.0627 | -104.1569
9.0663 | -104.1776
9.07 | -104.1985
9.0737 | -104.2191
9.0774 | -104.2406
9.0811 | -104.2618
9.0848 | -104.2812
9.0885 | -104.2953
9.0922 | -104.3039
9.0959 | -104.3148
9.0995 | -104.3273
9.1032 | -104.3361
9.1069 | -104.3406
9.1106 | -104.3476
9.1143 | -104.3598
9.118 | -104.3724
9.1217 | -104.3842
9.1254 | -104.3951
9.1291 | -104.4008
9.1327 | -104.4111
9.1364 | -104.4233
9.1401 | -104.4324
9.1438 | -104.4438
9.1475 | -104.459
9.1512 | -104.4758
9.1549 | -104.4906
9.1586 | -104.5017
9.1622 | -104.5076
9.1659 | -104.5129
9.1696 | -104.5175
9.1733 | -104.5178
9.177 | -104.5122
9.1807 | -104.5086
9.1844 | -104.5021
9.1881 | -104.4954
9.1918 | -104.4857
9.1954 | -104.4763
9.1991 | -104.4684
9.2028 | -104.4594
9.2065 | -104.4481
9.2102 | -104.4322
9.2139 | -104.4175
9.2176 | -104.4065
9.2213 | -104.3972
9.225 | -104.3928
9.2286 | -104.393
9.2323 | -104.3884
9.236 | -104.3817
9.2397 | -104.3785
9.2434 | -104.3735
9.2471 | -104.3625
9.2508 | -104.3489
9.2545 | -104.3371
9.2581 | -104.3284
9.2618 | -104.3182
9.2655 | -104.3075
9.2692 | -104.2935
9.2729 | -104.2822
9.2766 | -104.2715
9.2803 | -104.26
9.284 | -104.2502
9.2877 | -104.241
9.2913 | -104.2387
9.295 | -104.2412
9.2987 | -104.2451
9.3024 | -104.2494
9.3061 | -104.2567
9.3098 | -104.267
9.3135 | -104.2752
9.3172 | -104.2841
9.3209 | -104.2911
9.3245 | -104.2936
9.3282 | -104.297
9.3319 | -104.3017
9.3356 | -104.3038
9.3393 | -104.2977
9.343 | -104.2926
9.3467 | -104.2887
9.3504 | -104.2742
9.3541 | -104.2506
9.3577 | -104.2289
9.3614 | -104.2048
9.3651 | -104.1775
9.3688 | -104.1507
9.3725 | -104.1234
9.3762 | -104.0922
9.3799 | -104.0646
9.3836 | -104.0478
9.3872 | -104.033
9.3909 | -104.0187
9.3946 | -104.0045
9.3983 | -104.0001
9.402 | -103.9997
9.4057 | -104.0011
9.4094 | -104.0027
9.4131 | -104.0059
9.4168 | -104.0117
9.4204 | -104.0172
9.4241 | -104.0225
9.4278 | -104.0259
9.4315 | -104.0283
9.4352 | -104.0302
9.4389 | -104.031
9.4426 | -104.0288
9.4463 | -104.0235
9.45 | -104.0225
9.4536 | -104.0195
9.4573 | -104.0163
9.461 | -104.0158
9.4647 | -104.0129
9.4684 | -104.0074
9.4721 | -104.0019
9.4758 | -104.0009
9.4795 | -103.9954
9.4831 | -103.9869
9.4868 | -103.9819
9.4905 | -103.9756
9.4942 | -103.967
9.4979 | -103.958
9.5016 | -103.9495
9.5053 | -103.9453
9.509 | -103.9443
9.5127 | -103.9439
9.5163 | -103.9455
9.52 | -103.9493
9.5237 | -103.9586
9.5274 | -103.9726
9.5311 | -103.9898
9.5348 | -104.0076
9.5385 | -104.0327
9.5422 | -104.0616
9.5459 | -104.091
9.5495 | -104.121
9.5532 | -104.1474
9.5569 | -104.1717
9.5606 | -104.1972
9.5643 | -104.2204
9.568 | -104.2456
9.5717 | -104.266
9.5754 | -104.2694
9.579 | -104.2821
9.5827 | -104.2931
9.5864 | -104.3054
9.5901 | -104.3108
9.5938 | -104.3119
9.5975 | -104.3107
9.6012 | -104.3067
9.6049 | -104.2997
9.6086 | -104.2952
9.6122 | -104.2892
9.6159 | -104.2798
9.6196 | -104.2696
9.6233 | -104.2593
9.627 | -104.2467
9.6307 | -104.2323
9.6344 | -104.213
9.6381 | -104.1932
9.6418 | -104.1742
9.6454 | -104.1541
9.6491 | -104.1337
9.6528 | -104.1168
9.6565 | -104.0999
9.6602 | -104.0835
9.6639 | -104.0704
9.6676 | -104.062
9.6713 | -104.0596
9.675 | -104.06
9.6786 | -104.0627
9.6823 | -104.0684
9.686 | -104.0808
9.6897 | -104.0985
9.6934 | -104.1211
9.6971 | -104.1473
9.7008 | -104.1735
9.7045 | -104.1961
9.7081 | -104.2168
9.7118 | -104.2421
9.7155 | -104.274
9.7192 | -104.3013
9.7229 | -104.3263
9.7266 | -104.3499
9.7303 | -104.3733
9.734 | -104.398
9.7377 | -104.419
9.7413 | -104.4408
9.745 | -104.4638
9.7487 | -104.4882
9.7524 | -104.5087
9.7561 | -104.5254
9.7598 | -104.5409
9.7635 | -104.5557
9.7672 | -104.5714
9.7709 | -104.5858
9.7745 | -104.6026
9.7782 | -104.6187
9.7819 | -104.6341
9.7856 | -104.6461
9.7893 | -104.6484
9.793 | -104.6491
9.7967 | -104.6471
9.8004 | -104.6467
9.804 | -104.6445
9.8077 | -104.6406
9.8114 | -104.6321
9.8151 | -104.6191
9.8188 | -104.6023
9.8225 | -104.5825
9.8262 | -104.5627
9.8299 | -104.5396
9.8336 | -104.5118
9.8372 | -104.4837
9.8409 | -104.4609
9.8446 | -104.4405
9.8483 | -104.4272
9.852 | -104.415
9.8557 | -104.3997
9.8594 | -104.3857
9.8631 | -104.3793
9.8668 | -104.373
9.8704 | -104.3731
9.8741 | -104.3753
9.8778 | -104.3775
9.8815 | -104.3814
9.8852 | -104.3851
9.8889 | -104.3938
9.8926 | -104.4034
9.8963 | -104.4127
9.8999 | -104.4283
9.9036 | -104.4439
9.9073 | -104.4538
9.911 | -104.4659
9.9147 | -104.4782
9.9184 | -104.4862
9.9221 | -104.4989
9.9258 | -104.5105
9.9295 | -104.5177
9.9331 | -104.5323
9.9368 | -104.5419
9.9405 | -104.5459
9.9442 | -104.5443
9.9479 | -104.5378
9.9516 | -104.5287
9.9553 | -104.519
9.959 | -104.5082
9.9627 | -104.4993
9.9663 | -104.4852
9.97 | -104.4652
9.9737 | -104.4497
9.9774 | -104.4371
9.9811 | -104.4246
9.9848 | -104.4098
9.9885 | -104.4013
9.9922 | -104.3944
9.9959 | -104.39
9.9995 | -104.395
10.0032 | -104.4039
10.0069 | -104.4179
10.0106 | -104.4396
10.0143 | -104.4605
10.018 | -104.4873
10.0217 | -104.5184
10.0254 | -104.5535
10.029 | -104.5888
10.0327 | -104.6214
10.0364 | -104.6569
10.0401 | -104.6938
10.0438 | -104.7266
10.0475 | -104.7583
10.0512 | -104.7881
10.0549 | -104.8102
10.0586 | -104.8237
10.0622 | -104.8348
10.0659 | -104.842
10.0696 | -104.847
10.0733 | -104.8516
10.077 | -104.8473
10.0807 | -104.8416
10.0844 | -104.8354
10.0881 | -104.8268
10.0918 | -104.8159
10.0954 | -104.802
10.0991 | -104.7879
10.1028 | -104.7758
10.1065 | -104.762
10.1102 | -104.7481
10.1139 | -104.7349
10.1176 | -104.7279
10.1213 | -104.7185
10.1249 | -104.7077
10.1286 | -104.6976
10.1323 | -104.6917
10.136 | -104.6844
10.1397 | -104.6804
10.1434 | -104.6803
10.1471 | -104.6799
10.1508 | -104.6775
10.1545 | -104.6772
10.1581 | -104.6796
10.1618 | -104.6856
10.1655 | -104.6909
10.1692 | -104.6972
10.1729 | -104.7034
10.1766 | -104.715
10.1803 | -104.7246
10.184 | -104.735
10.1877 | -104.7465
10.1913 | -104.7576
10.195 | -104.7658
10.1987 | -104.7739
10.2024 | -104.7821
10.2061 | -104.788
10.2098 | -104.7853
10.2135 | -104.7779
10.2172 | -104.7625
10.2208 | -104.7465
10.2245 | -104.73
10.2282 | -104.7079
10.2319 | -104.6852
10.2356 | -104.6627
10.2393 | -104.6398
10.243 | -104.6182
10.2467 | -104.5975
10.2504 | -104.5781
10.254 | -104.5586
10.2577 | -104.5449
10.2614 | -104.5355
10.2651 | -104.5293
10.2688 | -104.5241
10.2725 | -104.5199
10.2762 | -104.5247
10.2799 | -104.5369
10.2836 | -104.5477
10.2872 | -104.5602
10.2909 | -104.5782
10.2946 | -104.5962
10.2983 | -104.6173
10.302 | -104.6408
10.3057 | -104.6638
10.3094 | -104.6838
10.3131 | -104.7016
10.3168 | -104.7191
10.3204 | -104.7349
10.3241 | -104.7508
10.3278 | -104.7591
10.3315 | -104.7674
10.3352 | -104.7745
10.3389 | -104.7765
10.3426 | -104.7788
10.3463 | -104.7775
10.3499 | -104.7727
10.3536 | -104.7602
10.3573 | -104.7491
10.361 | -104.7408
10.3647 | -104.729
10.3684 | -104.7235
10.3721 | -104.7169
10.3758 | -104.7068
10.3795 | -104.6988
10.3831 | -104.693
10.3868 | -104.6865
10.3905 | -104.6834
10.3942 | -104.6799
10.3979 | -104.6768
10.4016 | -104.6761
10.4053 | -104.68
10.409 | -104.6844
10.4127 | -104.6889
10.4163 | -104.6968
10.42 | -104.7075
10.4237 | -104.723
10.4274 | -104.7381
10.4311 | -104.7517
10.4348 | -104.7695
10.4385 | -104.7928
10.4422 | -104.8158
10.4458 | -104.8377
10.4495 | -104.8622
10.4532 | -104.8908
10.4569 | -104.9177
10.4606 | -104.9422
10.4643 | -104.9687
10.468 | -104.9941
10.4717 | -105.017
10.4754 | -105.037
10.479 | -105.057
10.4827 | -105.0734
10.4864 | -105.083
10.4901 | -105.0906
10.4938 | -105.0928
10.4975 | -105.0946
10.5012 | -105.0926
10.5049 | -105.0845
10.5086 | -105.07
10.5122 | -105.049
10.5159 | -105.0255
10.5196 | -105.0009
10.5233 | -104.9727
10.527 | -104.9444
10.5307 | -104.9176
10.5344 | -104.8843
10.5381 | -104.8561
10.5417 | -104.8321
10.5454 | -104.809
10.5491 | -104.784
10.5528 | -104.7613
10.5565 | -104.7384
10.5602 | -104.7136
10.5639 | -104.6904
10.5676 | -104.6675
10.5713 | -104.6492
10.5749 | -104.634
10.5786 | -104.6184
10.5823 | -104.6049
10.586 | -104.5894
10.5897 | -104.5754
10.5934 | -104.5637
10.5971 | -104.557
10.6008 | -104.5498
10.6045 | -104.5422
10.6081 | -104.54
10.6118 | -104.5448
10.6155 | -104.5454
10.6192 | -104.5478
10.6229 | -104.5531
10.6266 | -104.5588
10.6303 | -104.566
10.634 | -104.5757
10.6377 | -104.5884
10.6413 | -104.6123
10.645 | -104.6308
10.6487 | -104.651
10.6524 | -104.676
10.6561 | -104.6994
10.6598 | -104.7266
10.6635 | -104.7572
10.6672 | -104.7877
10.6708 | -104.8167
10.6745 | -104.8411
10.6782 | -104.8655
10.6819 | -104.8902
10.6856 | -104.9096
10.6893 | -104.9281
10.693 | -104.9405
10.6967 | -104.9468
10.7004 | -104.9498
10.704 | -104.9473
10.7077 | -104.9408
10.7114 | -104.9342
10.7151 | -104.9256
10.7188 | -104.9124
10.7225 | -104.8973
10.7262 | -104.8798
10.7299 | -104.8589
10.7336 | -104.8395
10.7372 | -104.823
10.7409 | -104.8106
10.7446 | -104.7931
10.7483 | -104.7752
10.752 | -104.7588
10.7557 | -104.7445
10.7594 | -104.7312
10.7631 | -104.718
10.7667 | -104.7062
10.7704 | -104.6999
10.7741 | -104.6959
10.7778 | -104.6983
10.7815 | -104.7016
10.7852 | -104.7075
10.7889 | -104.7193
10.7926 | -104.7347
10.7963 | -104.7566
10.7999 | -104.7815
10.8036 | -104.808
10.8073 | -104.8359
10.811 | -104.8635
10.8147 | -104.8943
10.8184 | -104.9236
10.8221 | -104.9465
10.8258 | -104.9683
10.8295 | -104.9834
10.8331 | -104.9972
10.8368 | -105.0088
10.8405 | -105.0149
10.8442 | -105.0141
10.8479 | -105.0087
10.8516 | -104.9991
10.8553 | -104.9849
10.859 | -104.9675
10.8626 | -104.9455
10.8663 | -104.9185
10.87 | -104.8897
10.8737 | -104.8627
10.8774 | -104.834
10.8811 | -104.8023
10.8848 | -104.7706
10.8885 | -104.746
10.8922 | -104.7241
10.8958 | -104.7034
10.8995 | -104.6882
10.9032 | -104.6764
10.9069 | -104.669
10.9106 | -104.6679
10.9143 | -104.6702
10.918 | -104.6809
10.9217 | -104.6957
10.9254 | -104.711
10.929 | -104.7245
10.9327 | -104.7372
10.9364 | -104.753
10.9401 | -104.771
10.9438 | -104.7902
10.9475 | -104.8016
10.9512 | -104.8097
10.9549 | -104.8123
10.9586 | -104.8135
10.9622 | -104.8117
10.9659 | -104.8067
10.9696 | -104.7977
10.9733 | -104.7904
10.977 | -104.7845
10.9807 | -104.7763
10.9844 | -104.7694
10.9881 | -104.7677
10.9917 | -104.766
10.9954 | -104.7652
10.9991 | -104.768
11.0028 | -104.7752
11.0065 | -104.7857
11.0102 | -104.7985
11.0139 | -104.8139
11.0176 | -104.8336
11.0213 | -104.8582
11.0249 | -104.8826
11.0286 | -104.9051
11.0323 | -104.9259
11.036 | -104.9439
11.0397 | -104.9641
11.0434 | -104.9837
11.0471 | -104.994
11.0508 | -105.0041
11.0545 | -105.0141
11.0581 | -105.0208
11.0618 | -105.0294
11.0655 | -105.0345
11.0692 | -105.0362
11.0729 | -105.0359
11.0766 | -105.0326
11.0803 | -105.0311
11.084 | -105.0344
11.0876 | -105.0401
11.0913 | -105.0485
11.095 | -105.0524
11.0987 | -105.0542
11.1024 | -105.0573
11.1061 | -105.0672
11.1098 | -105.0774
11.1135 | -105.0888
11.1172 | -105.1012
11.1208 | -105.112
11.1245 | -105.1241
11.1282 | -105.1388
11.1319 | -105.1464
11.1356 | -105.1479
11.1393 | -105.1523
11.143 | -105.155
11.1467 | -105.1542
11.1504 | -105.1519
11.154 | -105.1482
11.1577 | -105.14
11.1614 | -105.1332
11.1651 | -105.1284
11.1688 | -105.1237
11.1725 | -105.1229
11.1762 | -105.1204
11.1799 | -105.119
11.1835 | -105.1172
11.1872 | -105.118
11.1909 | -105.1225
11.1946 | -105.1294
11.1983 | -105.141
11.202 | -105.1534
11.2057 | -105.1626
11.2094 | -105.1746
11.2131 | -105.1874
11.2167 | -105.1949
11.2204 | -105.2019
11.2241 | -105.2107
11.2278 | -105.2174
11.2315 | -105.2236
11.2352 | -105.2309
11.2389 | -105.2349
11.2426 | -105.2371
11.2463 | -105.2334
11.2499 | -105.2221
11.2536 | -105.21
11.2573 | -105.1974
11.261 | -105.1823
11.2647 | -105.1638
11.2684 | -105.1454
11.2721 | -105.1249
11.2758 | -105.1011
11.2795 | -105.078
11.2831 | -105.0546
11.2868 | -105.0292
11.2905 | -105.0042
11.2942 | -104.9837
11.2979 | -104.9632
11.3016 | -104.9402
11.3053 | -104.9167
11.309 | -104.8941
11.3126 | -104.8692
11.3163 | -104.8429
11.32 | -104.8222
11.3237 | -104.8012
11.3274 | -104.7767
11.3311 | -104.7525
11.3348 | -104.7281
11.3385 | -104.706
11.3422 | -104.6823
11.3458 | -104.6586
11.3495 | -104.6361
11.3532 | -104.6119
11.3569 | -104.5889
11.3606 | -104.5684
11.3643 | -104.5514
11.368 | -104.535
11.3717 | -104.5234
11.3754 | -104.516
11.379 | -104.5081
11.3827 | -104.5031
11.3864 | -104.5007
11.3901 | -104.4979
11.3938 | -104.4979
11.3975 | -104.5001
11.4012 | -104.5033
11.4049 | -104.51
11.4085 | -104.5193
11.4122 | -104.5324
11.4159 | -104.5491
11.4196 | -104.5607
11.4233 | -104.572
11.427 | -104.5893
11.4307 | -104.6068
11.4344 | -104.6229
11.4381 | -104.6391
11.4417 | -104.653
11.4454 | -104.6662
11.4491 | -104.6826
11.4528 | -104.6976
11.4565 | -104.7081
11.4602 | -104.7161
11.4639 | -104.7253
11.4676 | -104.7361
11.4713 | -104.7449
11.4749 | -104.7531
11.4786 | -104.7619
11.4823 | -104.7714
11.486 | -104.7749
11.4897 | -104.7764
11.4934 | -104.7793
11.4971 | -104.7816
11.5008 | -104.7842
11.5044 | -104.787
11.5081 | -104.7887
11.5118 | -104.791
11.5155 | -104.793
11.5192 | -104.7967
11.5229 | -104.7991
11.5266 | -104.8
11.5303 | -104.8004
11.534 | -104.8018
11.5376 | -104.803
11.5413 | -104.806
11.545 | -104.8086
11.5487 | -104.812
11.5524 | -104.8165
11.5561 | -104.819
11.5598 | -104.8206
11.5635 | -104.8196
11.5672 | -104.8181
11.5708 | -104.8152
11.5745 | -104.8131
11.5782 | -104.8094
11.5819 | -104.8095
11.5856 | -104.8114
11.5893 | -104.8107
11.593 | -104.8054
11.5967 | -104.8012
11.6004 | -104.7993
11.604 | -104.7997
11.6077 | -104.8023
11.6114 | -104.8048
11.6151 | -104.8074
11.6188 | -104.8105
11.6225 | -104.8147
11.6262 | -104.8175
11.6299 | -104.821
11.6335 | -104.823
11.6372 | -104.8272
11.6409 | -104.8307
11.6446 | -104.8351
11.6483 | -104.8392
11.652 | -104.8442
11.6557 | -104.8504
11.6594 | -104.8567
11.6631 | -104.8619
11.6667 | -104.8683
11.6704 | -104.874
11.6741 | -104.8757
11.6778 | -104.8734
11.6815 | -104.8705
11.6852 | -104.8687
11.6889 | -104.8641
11.6926 | -104.8596
11.6963 | -104.8535
11.6999 | -104.8506
11.7036 | -104.8537
11.7073 | -104.8509
11.711 | -104.8506
11.7147 | -104.8547
11.7184 | -104.8585
11.7221 | -104.8625
11.7258 | -104.8719
11.7294 | -104.8832
11.7331 | -104.896
11.7368 | -104.9106
11.7405 | -104.9282
11.7442 | -104.9452
11.7479 | -104.9662
11.7516 | -104.9867
11.7553 | -105.0098
11.759 | -105.0368
11.7626 | -105.0647
11.7663 | -105.0892
11.77 | -105.1121
11.7737 | -105.1347
11.7774 | -105.1573
11.7811 | -105.1817
11.7848 | -105.2031
11.7885 | -105.2246
11.7922 | -105.2463
11.7958 | -105.2643
11.7995 | -105.2803
11.8032 | -105.2956
11.8069 | -105.3026
11.8106 | -105.3099
11.8143 | -105.3177
11.818 | -105.3199
11.8217 | -105.3185
11.8253 | -105.3128
11.829 | -105.3043
11.8327 | -105.2956
11.8364 | -105.2821
11.8401 | -105.2662
11.8438 | -105.2524
11.8475 | -105.2366
11.8512 | -105.2188
11.8549 | -105.1975
11.8585 | -105.1773
11.8622 | -105.1569
11.8659 | -105.1348
11.8696 | -105.1104
11.8733 | -105.0861
11.877 | -105.0657
11.8807 | -105.0494
11.8844 | -105.0346
11.8881 | -105.017
11.8917 | -105.006
11.8954 | -104.9982
11.8991 | -104.9857
11.9028 | -104.9783
11.9065 | -104.9764
11.9102 | -104.976
11.9139 | -104.9731
11.9176 | -104.9719
11.9213 | -104.9712
11.9249 | -104.9688
11.9286 | -104.9619
11.9323 | -104.9536
11.936 | -104.949
11.9397 | -104.9449
11.9434 | -104.9363
11.9471 | -104.9285
11.9508 | -104.9203
11.9544 | -104.9042
11.9581 | -104.886
11.9618 | -104.8692
11.9655 | -104.8551
11.9692 | -104.8423
11.9729 | -104.8223
11.9766 | -104.7994
11.9803 | -104.784
11.984 | -104.7669
11.9876 | -104.7457
11.9913 | -104.7312
11.995 | -104.7197
11.9987 | -104.7094
12.0024 | -104.6984
12.0061 | -104.6862
12.0098 | -104.6778
12.0135 | -104.6707
12.0172 | -104.6623
12.0208 | -104.6575
12.0245 | -104.6546
12.0282 | -104.6537
12.0319 | -104.6536
12.0356 | -104.6581
12.0393 | -104.6641
12.043 | -104.6685
12.0467 | -104.6742
12.0503 | -104.681
12.054 | -104.6886
12.0577 | -104.697
12.0614 | -104.7062
12.0651 | -104.7141
12.0688 | -104.7234
12.0725 | -104.7384
12.0762 | -104.7503
12.0799 | -104.759
12.0835 | -104.7707
12.0872 | -104.7881
12.0909 | -104.8048
12.0946 | -104.8229
12.0983 | -104.8458
12.102 | -104.8729
12.1057 | -104.9012
12.1094 | -104.9274
12.1131 | -104.955
12.1167 | -104.9869
12.1204 | -105.0158
12.1241 | -105.0448
12.1278 | -105.0823
12.1315 | -105.1225
12.1352 | -105.1584
12.1389 | -105.1917
12.1426 | -105.2296
12.1462 | -105.2707
12.1499 | -105.3128
12.1536 | -105.3514
12.1573 | -105.3897
12.161 | -105.4258
12.1647 | -105.4551
12.1684 | -105.4796
12.1721 | -105.5032
12.1758 | -105.5279
12.1794 | -105.5507
12.1831 | -105.566
12.1868 | -105.5791
12.1905 | -105.5896
12.1942 | -105.5921
12.1979 | -105.5889
12.2016 | -105.5825
12.2053 | -105.5704
12.209 | -105.5495
12.2126 | -105.5279
12.2163 | -105.5054
12.22 | -105.4791
12.2237 | -105.447
12.2274 | -105.4086
12.2311 | -105.3662
12.2348 | -105.3276
12.2385 | -105.2751
12.2422 | -105.2325
12.2458 | -105.1896
12.2495 | -105.1466
12.2532 | -105.1018
12.2569 | -105.0572
12.2606 | -105.0138
12.2643 | -104.9732
12.268 | -104.9357
12.2717 | -104.898
12.2753 | -104.8621
12.279 | -104.8288
12.2827 | -104.8002
12.2864 | -104.7753
12.2901 | -104.7537
12.2938 | -104.7355
12.2975 | -104.7191
12.3012 | -104.7021
12.3049 | -104.6877
12.3085 | -104.6759
12.3122 | -104.6694
12.3159 | -104.6646
12.3196 | -104.661
12.3233 | -104.661
12.327 | -104.6636
12.3307 | -104.6679
12.3344 | -104.6772
12.3381 | -104.6927
12.3417 | -104.7124
12.3454 | -104.7343
12.3491 | -104.7571
12.3528 | -104.7797
12.3565 | -104.8065
12.3602 | -104.8394
12.3639 | -104.874
12.3676 | -104.91
12.3712 | -104.9476
12.3749 | -104.9854
12.3786 | -105.0228
12.3823 | -105.0625
12.386 | -105.1024
12.3897 | -105.1404
12.3934 | -105.1782
12.3971 | -105.2125
12.4008 | -105.2475
12.4044 | -105.2836
12.4081 | -105.3156
12.4118 | -105.3442
12.4155 | -105.3678
12.4192 | -105.3889
12.4229 | -105.4097
12.4266 | -105.4258
12.4303 | -105.4378
12.434 | -105.4499
12.4376 | -105.4599
12.4413 | -105.4671
12.445 | -105.4675
12.4487 | -105.4665
12.4524 | -105.4679
12.4561 | -105.4674
12.4598 | -105.4687
12.4635 | -105.4663
12.4671 | -105.4652
12.4708 | -105.4654
12.4745 | -105.4665
12.4782 | -105.4706
12.4819 | -105.4782
12.4856 | -105.4808
12.4893 | -105.4808
12.493 | -105.4832
12.4967 | -105.4869
12.5003 | -105.4924
12.504 | -105.5013
12.5077 | -105.5086
12.5114 | -105.5155
12.5151 | -105.5226
12.5188 | -105.5303
12.5225 | -105.5358
12.5262 | -105.5438
12.5299 | -105.5513
12.5335 | -105.5609
12.5372 | -105.567
12.5409 | -105.5673
12.5446 | -105.5656
12.5483 | -105.5664
12.552 | -105.5643
12.5557 | -105.5604
12.5594 | -105.5567
12.5631 | -105.5493
12.5667 | -105.5421
12.5704 | -105.5379
12.5741 | -105.5309
12.5778 | -105.5256
12.5815 | -105.5184
12.5852 | -105.5103
12.5889 | -105.5025
12.5926 | -105.4953
12.5962 | -105.4914
12.5999 | -105.486
12.6036 | -105.4773
12.6073 | -105.4679
12.611 | -105.4567
12.6147 | -105.4464
12.6184 | -105.4373
12.6221 | -105.4256
12.6258 | -105.4156
12.6294 | -105.4053
12.6331 | -105.3933
12.6368 | -105.3792
12.6405 | -105.364
12.6442 | -105.3466
12.6479 | -105.3279
12.6516 | -105.3104
12.6553 | -105.2939
12.659 | -105.2781
12.6626 | -105.2597
12.6663 | -105.246
12.67 | -105.2378
12.6737 | -105.2269
12.6774 | -105.217
12.6811 | -105.2075
12.6848 | -105.1999
12.6885 | -105.1944
12.6921 | -105.1907
12.6958 | -105.188
12.6995 | -105.1866
12.7032 | -105.1894
12.7069 | -105.1921
12.7106 | -105.1978
12.7143 | -105.2054
12.718 | -105.2174
12.7217 | -105.2332
12.7253 | -105.2484
12.729 | -105.2642
12.7327 | -105.285
12.7364 | -105.3031
12.7401 | -105.319
12.7438 | -105.3374
12.7475 | -105.358
12.7512 | -105.3729
12.7549 | -105.3891
12.7585 | -105.4061
12.7622 | -105.4231
12.7659 | -105.4382
12.7696 | -105.4429
12.7733 | -105.4566
12.777 | -105.4651
12.7807 | -105.4724
12.7844 | -105.4772
12.788 | -105.4789
12.7917 | -105.4781
12.7954 | -105.4795
12.7991 | -105.4774
12.8028 | -105.4693
12.8065 | -105.4591
12.8102 | -105.4472
12.8139 | -105.4356
12.8176 | -105.424
12.8212 | -105.4107
12.8249 | -105.3959
12.8286 | -105.3812
12.8323 | -105.3667
12.836 | -105.3519
12.8397 | -105.338
12.8434 | -105.3235
12.8471 | -105.3117
12.8508 | -105.2985
12.8544 | -105.2834
12.8581 | -105.267
12.8618 | -105.2521
12.8655 | -105.2403
12.8692 | -105.2289
12.8729 | -105.2207
12.8766 | -105.2139
12.8803 | -105.206
12.884 | -105.1976
12.8876 | -105.1901
12.8913 | -105.1852
12.895 | -105.1806
12.8987 | -105.1758
12.9024 | -105.1718
12.9061 | -105.1694
12.9098 | -105.1691
12.9135 | -105.169
12.9171 | -105.1694
12.9208 | -105.1747
12.9245 | -105.1801
12.9282 | -105.1893
12.9319 | -105.1955
12.9356 | -105.2034
12.9393 | -105.2105
12.943 | -105.2182
12.9467 | -105.2264
12.9503 | -105.2375
12.954 | -105.2513
12.9577 | -105.2637
12.9614 | -105.2744
12.9651 | -105.2855
12.9688 | -105.2979
12.9725 | -105.3134
12.9762 | -105.3297
12.9799 | -105.3464
12.9835 | -105.3632
12.9872 | -105.3764
12.9909 | -105.3916
12.9946 | -105.405
12.9983 | -105.419
13.002 | -105.4311
13.0057 | -105.4412
13.0094 | -105.4513
13.013 | -105.4585
13.0167 | -105.4669
13.0204 | -105.4706
13.0241 | -105.4705
13.0278 | -105.4637
13.0315 | -105.4595
13.0352 | -105.4554
13.0389 | -105.4472
13.0426 | -105.4374
13.0462 | -105.4298
13.0499 | -105.4198
13.0536 | -105.4053
13.0573 | -105.3893
13.061 | -105.3736
13.0647 | -105.3601
13.0684 | -105.3475
13.0721 | -105.3364
13.0758 | -105.3254
13.0794 | -105.3139
13.0831 | -105.3033
13.0868 | -105.2938
13.0905 | -105.2858
13.0942 | -105.2816
13.0979 | -105.2789
13.1016 | -105.2784
13.1053 | -105.2786
13.1089 | -105.2811
13.1126 | -105.2872
13.1163 | -105.2943
13.12 | -105.3053
13.1237 | -105.3181
13.1274 | -105.3328
13.1311 | -105.349
13.1348 | -105.3694
13.1385 | -105.388
13.1421 | -105.406
13.1458 | -105.4277
13.1495 | -105.4507
13.1532 | -105.4761
13.1569 | -105.4991
13.1606 | -105.5192
13.1643 | -105.5401
13.168 | -105.5605
13.1717 | -105.5778
13.1753 | -105.599
13.179 | -105.6173
13.1827 | -105.6368
13.1864 | -105.6544
13.1901 | -105.6703
13.1938 | -105.6843
13.1975 | -105.6992
13.2012 | -105.7147
13.2049 | -105.7275
13.2085 | -105.7406
13.2122 | -105.7541
13.2159 | -105.7695
13.2196 | -105.7786
13.2233 | -105.783
13.227 | -105.7893
13.2307 | -105.7953
13.2344 | -105.8005
13.238 | -105.8052
13.2417 | -105.811
13.2454 | -105.8119
13.2491 | -105.8101
13.2528 | -105.8048
13.2565 | -105.8006
13.2602 | -105.7972
13.2639 | -105.7915
13.2676 | -105.7845
13.2712 | -105.7751
13.2749 | -105.7655
13.2786 | -105.7549
13.2823 | -105.7396
13.286 | -105.7225
13.2897 | -105.7042
13.2934 | -105.6882
13.2971 | -105.6688
13.3008 | -105.639
13.3044 | -105.6142
13.3081 | -105.5907
13.3118 | -105.5655
13.3155 | -105.5416
13.3192 | -105.5162
13.3229 | -105.4908
13.3266 | -105.4632
13.3303 | -105.4408
13.3339 | -105.4189
13.3376 | -105.3956
13.3413 | -105.373
13.345 | -105.3538
13.3487 | -105.3377
13.3524 | -105.3231
13.3561 | -105.3111
13.3598 | -105.3004
13.3635 | -105.294
13.3671 | -105.2888
13.3708 | -105.287
13.3745 | -105.2877
13.3782 | -105.2888
13.3819 | -105.2911
13.3856 | -105.2973
13.3893 | -105.3062
13.393 | -105.3141
13.3967 | -105.3226
13.4003 | -105.3345
13.404 | -105.3484
13.4077 | -105.3607
13.4114 | -105.3733
13.4151 | -105.386
13.4188 | -105.4009
13.4225 | -105.4152
13.4262 | -105.4257
13.4298 | -105.4343
13.4335 | -105.4428
13.4372 | -105.4538
13.4409 | -105.463
13.4446 | -105.4713
13.4483 | -105.4733
13.452 | -105.4724
13.4557 | -105.4736
13.4594 | -105.4745
13.463 | -105.4747
13.4667 | -105.474
13.4704 | -105.4711
13.4741 | -105.4712
13.4778 | -105.4735
13.4815 | -105.4707
13.4852 | -105.4657
13.4889 | -105.4609
13.4926 | -105.4549
13.4962 | -105.4509
13.4999 | -105.446
13.5036 | -105.4392
13.5073 | -105.433
13.511 | -105.4269
13.5147 | -105.4211
13.5184 | -105.4144
13.5221 | -105.4078
13.5258 | -105.4013
13.5294 | -105.3966
13.5331 | -105.3923
13.5368 | -105.3902
13.5405 | -105.3859
13.5442 | -105.3845
13.5479 | -105.386
13.5516 | -105.3883
13.5553 | -105.394
13.5589 | -105.3998
13.5626 | -105.4083
13.5663 | -105.4198
13.57 | -105.431
13.5737 | -105.4428
13.5774 | -105.4561
13.5811 | -105.47
13.5848 | -105.484
13.5885 | -105.4999
13.5921 | -105.5133
13.5958 | -105.5245
13.5995 | -105.5336
13.6032 | -105.5417
13.6069 | -105.5498
13.6106 | -105.5588
13.6143 | -105.5645
13.618 | -105.5679
13.6217 | -105.572
13.6253 | -105.5744
13.629 | -105.5748
13.6327 | -105.572
13.6364 | -105.569
13.6401 | -105.5646
13.6438 | -105.5614
13.6475 | -105.5551
13.6512 | -105.5474
13.6548 | -105.5386
13.6585 | -105.5271
13.6622 | -105.5144
13.6659 | -105.5034
13.6696 | -105.4913
13.6733 | -105.4788
13.677 | -105.4695
13.6807 | -105.4597
13.6844 | -105.4485
13.688 | -105.4391
13.6917 | -105.4311
13.6954 | -105.4233
13.6991 | -105.4138
13.7028 | -105.4072
13.7065 | -105.4004
13.7102 | -105.3947
13.7139 | -105.3877
13.7176 | -105.3815
13.7212 | -105.3738
13.7249 | -105.3677
13.7286 | -105.3657
13.7323 | -105.36
13.736 | -105.3554
13.7397 | -105.3492
13.7434 | -105.3416
13.7471 | -105.3336
13.7507 | -105.3284
13.7544 | -105.3226
13.7581 | -105.3127
13.7618 | -105.3037
13.7655 | -105.2979
13.7692 | -105.2934
13.7729 | -105.288
13.7766 | -105.2812
13.7803 | -105.2771
13.7839 | -105.2738
13.7876 | -105.2721
13.7913 | -105.2683
13.795 | -105.2655
13.7987 | -105.2648
13.8024 | -105.2646
13.8061 | -105.2631
13.8098 | -105.2591
13.8135 | -105.2578
13.8171 | -105.2581
13.8208 | -105.26
13.8245 | -105.2633
13.8282 | -105.2698
13.8319 | -105.2826
13.8356 | -105.2907
13.8393 | -105.2997
13.843 | -105.3125
13.8467 | -105.3271
13.8503 | -105.3407
13.854 | -105.3569
13.8577 | -105.3752
13.8614 | -105.3922
13.8651 | -105.4096
13.8688 | -105.4275
13.8725 | -105.4459
13.8762 | -105.4602
13.8798 | -105.4743
13.8835 | -105.4925
13.8872 | -105.508
13.8909 | -105.5204
13.8946 | -105.5316
13.8983 | -105.5453
13.902 | -105.5577
13.9057 | -105.567
13.9094 | -105.5723
13.913 | -105.578
13.9167 | -105.5838
13.9204 | -105.5877
13.9241 | -105.5878
13.9278 | -105.5843
13.9315 | -105.5844
13.9352 | -105.5805
13.9389 | -105.574
13.9426 | -105.5684
13.9462 | -105.5622
13.9499 | -105.5516
13.9536 | -105.5423
13.9573 | -105.5273
13.961 | -105.4857
13.9647 | -105.4432
13.9684 | -105.416
13.9721 | -105.3906
13.9757 | -105.3639
13.9794 | -105.3322
13.9831 | -105.2984
13.9868 | -105.264
13.9905 | -105.2312
13.9942 | -105.1945
13.9979 | -105.1555
14.0016 | -105.1161
14.0053 | -105.0747
14.0089 | -105.0327
14.0126 | -104.992
14.0163 | -104.9498
14.02 | -104.9088
14.0237 | -104.8708
14.0274 | -104.8362
14.0311 | -104.8032
14.0348 | -104.7749
14.0385 | -104.7502
14.0421 | -104.7295
14.0458 | -104.7137
14.0495 | -104.7022
14.0532 | -104.6952
14.0569 | -104.6932
14.0606 | -104.6979
14.0643 | -104.7074
14.068 | -104.7208
14.0716 | -104.739
14.0753 | -104.7621
14.079 | -104.7918
14.0827 | -104.8263
14.0864 | -104.8611
14.0901 | -104.8975
14.0938 | -104.9355
14.0975 | -104.9744
14.1012 | -105.0122
14.1048 | -105.048
14.1085 | -105.0873
14.1122 | -105.1273
14.1159 | -105.1627
14.1196 | -105.193
14.1233 | -105.2229
14.127 | -105.2519
14.1307 | -105.2782
14.1344 | -105.2995
14.138 | -105.3181
14.1417 | -105.3341
14.1454 | -105.3469
14.1491 | -105.3575
14.1528 | -105.3685
14.1565 | -105.3779
14.1602 | -105.3919
14.1639 | -105.4214
14.1676 | -105.4379
14.1712 | -105.437
14.1749 | -105.4341
14.1786 | -105.4333
14.1823 | -105.4346
14.186 | -105.4391
14.1897 | -105.4437
14.1934 | -105.4466
14.1971 | -105.4534
14.2007 | -105.4636
14.2044 | -105.4735
14.2081 | -105.4855
14.2118 | -105.4971
14.2155 | -105.5086
14.2192 | -105.5233
14.2229 | -105.5355
14.2266 | -105.5453
14.2303 | -105.5569
14.2339 | -105.5683
14.2376 | -105.578
14.2413 | -105.5889
14.245 | -105.6001
14.2487 | -105.6083
14.2524 | -105.6187
14.2561 | -105.6301
14.2598 | -105.6415
14.2635 | -105.6496
14.2671 | -105.6579
14.2708 | -105.6656
14.2745 | -105.6693
14.2782 | -105.671
14.2819 | -105.6705
14.2856 | -105.6642
14.2893 | -105.6595
14.293 | -105.6565
14.2966 | -105.6502
14.3003 | -105.643
14.304 | -105.6353
14.3077 | -105.6298
14.3114 | -105.62
14.3151 | -105.6085
14.3188 | -105.5994
14.3225 | -105.5913
14.3262 | -105.581
14.3298 | -105.569
14.3335 | -105.5571
14.3372 | -105.5454
14.3409 | -105.536
14.3446 | -105.5242
14.3483 | -105.5098
14.352 | -105.497
14.3557 | -105.4842
14.3594 | -105.4707
14.363 | -105.4588
14.3667 | -105.4484
14.3704 | -105.4361
14.3741 | -105.428
14.3778 | -105.4192
14.3815 | -105.4088
14.3852 | -105.3983
14.3889 | -105.3908
14.3925 | -105.3858
14.3962 | -105.3817
14.3999 | -105.3795
14.4036 | -105.3792
14.4073 | -105.3767
14.411 | -105.3725
14.4147 | -105.3709
14.4184 | -105.3738
14.4221 | -105.3749
14.4257 | -105.375
14.4294 | -105.3783
14.4331 | -105.3822
14.4368 | -105.3857
14.4405 | -105.3931
14.4442 | -105.4022
14.4479 | -105.4076
14.4516 | -105.4118
14.4553 | -105.418
14.4589 | -105.4245
14.4626 | -105.4287
14.4663 | -105.4336
14.47 | -105.4419
14.4737 | -105.452
14.4774 | -105.4621
14.4811 | -105.4716
14.4848 | -105.4788
14.4885 | -105.4863
14.4921 | -105.496
14.4958 | -105.5039
14.4995 | -105.5098
14.5032 | -105.5155
14.5069 | -105.5221
14.5106 | -105.5249
14.5143 | -105.5282
14.518 | -105.5326
14.5216 | -105.5395
14.5253 | -105.5446
14.529 | -105.5476
14.5327 | -105.5493
14.5364 | -105.5544
14.5401 | -105.5558
14.5438 | -105.5609
14.5475 | -105.564
14.5512 | -105.5702
14.5548 | -105.5747
14.5585 | -105.5759
14.5622 | -105.5774
14.5659 | -105.5807
14.5696 | -105.5844
14.5733 | -105.5884
14.577 | -105.5902
14.5807 | -105.5926
14.5844 | -105.5935
14.588 | -105.5942
14.5917 | -105.5962
14.5954 | -105.5972
14.5991 | -105.5945
14.6028 | -105.5902
14.6065 | -105.5878
14.6102 | -105.5865
14.6139 | -105.5823
14.6175 | -105.5773
14.6212 | -105.5744
14.6249 | -105.5717
14.6286 | -105.5668
14.6323 | -105.5649
14.636 | -105.5631
14.6397 | -105.5608
14.6434 | -105.5595
14.6471 | -105.5579
14.6507 | -105.5559
14.6544 | -105.5551
14.6581 | -105.5551
14.6618 | -105.5562
14.6655 | -105.5575
14.6692 | -105.559
14.6729 | -105.5634
14.6766 | -105.566
14.6803 | -105.5666
14.6839 | -105.5673
14.6876 | -105.5668
14.6913 | -105.5694
14.695 | -105.5728
14.6987 | -105.5762
14.7024 | -105.5763
14.7061 | -105.5734
14.7098 | -105.5723
14.7134 | -105.5718
14.7171 | -105.5689
14.7208 | -105.5668
14.7245 | -105.5626
14.7282 | -105.5607
14.7319 | -105.5592
14.7356 | -105.5568
14.7393 | -105.5563
14.743 | -105.556
14.7466 | -105.5534
14.7503 | -105.5537
14.754 | -105.5499
14.7577 | -105.5471
14.7614 | -105.5432
14.7651 | -105.5412
14.7688 | -105.5409
14.7725 | -105.5402
14.7762 | -105.5388
14.7798 | -105.5373
14.7835 | -105.5372
14.7872 | -105.5383
14.7909 | -105.5372
14.7946 | -105.5381
14.7983 | -105.5388
14.802 | -105.5403
14.8057 | -105.5423
14.8094 | -105.5459
14.813 | -105.5503
14.8167 | -105.5558
14.8204 | -105.5588
14.8241 | -105.5627
14.8278 | -105.569
14.8315 | -105.5748
14.8352 | -105.5783
14.8389 | -105.5827
14.8425 | -105.5829
14.8462 | -105.5846
14.8499 | -105.5882
14.8536 | -105.5926
14.8573 | -105.5945
14.861 | -105.5977
14.8647 | -105.601
14.8684 | -105.6046
14.8721 | -105.6092
14.8757 | -105.6112
14.8794 | -105.6085
14.8831 | -105.6068
14.8868 | -105.6066
14.8905 | -105.6062
14.8942 | -105.6031
14.8979 | -105.5989
14.9016 | -105.5964
14.9053 | -105.5929
14.9089 | -105.5881
14.9126 | -105.5822
14.9163 | -105.5775
14.92 | -105.5711
14.9237 | -105.5629
14.9274 | -105.5536
14.9311 | -105.5434
14.9348 | -105.5328
14.9384 | -105.5235
14.9421 | -105.5118
14.9458 | -105.4961
14.9495 | -105.4829
14.9532 | -105.4666
14.9569 | -105.4521
14.9606 | -105.4387
14.9643 | -105.4256
14.968 | -105.4111
14.9716 | -105.3952
14.9753 | -105.3816
14.979 | -105.3677
14.9827 | -105.3508
14.9864 | -105.3317
14.9901 | -105.3124
14.9938 | -105.2956
14.9975 | -105.2836
15.0012 | -105.2733
15.0048 | -105.2634
15.0085 | -105.2509
15.0122 | -105.2412
15.0159 | -105.2326
15.0196 | -105.2264
15.0233 | -105.2218
15.027 | -105.216
15.0307 | -105.2125
15.0343 | -105.2132
15.038 | -105.2148
15.0417 | -105.2171
15.0454 | -105.2197
15.0491 | -105.2228
15.0528 | -105.2286
15.0565 | -105.2368
15.0602 | -105.2466
15.0639 | -105.2534
15.0675 | -105.2571
15.0712 | -105.2638
15.0749 | -105.2727
15.0786 | -105.2828
15.0823 | -105.2946
15.086 | -105.3055
15.0897 | -105.318
15.0934 | -105.3337
15.0971 | -105.3456
15.1007 | -105.3556
15.1044 | -105.3644
15.1081 | -105.3735
15.1118 | -105.3833
15.1155 | -105.3927
15.1192 | -105.3996
15.1229 | -105.4056
15.1266 | -105.4128
15.1303 | -105.418
15.1339 | -105.4207
15.1376 | -105.4247
15.1413 | -105.4264
15.145 | -105.4276
15.1487 | -105.4281
15.1524 | -105.4277
15.1561 | -105.4249
15.1598 | -105.4231
15.1634 | -105.4235
15.1671 | -105.4238
15.1708 | -105.4248
15.1745 | -105.4244
15.1782 | -105.4229
15.1819 | -105.4215
15.1856 | -105.4232
15.1893 | -105.4266
15.193 | -105.4289
15.1966 | -105.4336
15.2003 | -105.4384
15.204 | -105.4469
15.2077 | -105.4558
15.2114 | -105.4617
15.2151 | -105.4645
15.2188 | -105.4697
15.2225 | -105.4782
15.2262 | -105.487
15.2298 | -105.492
15.2335 | -105.4947
15.2372 | -105.4975
15.2409 | -105.5024
15.2446 | -105.5104
15.2483 | -105.5137
15.252 | -105.5147
15.2557 | -105.5142
15.2593 | -105.5158
15.263 | -105.5189
15.2667 | -105.5189
15.2704 | -105.5155
15.2741 | -105.5125
15.2778 | -105.5076
15.2815 | -105.5035
15.2852 | -105.499
15.2889 | -105.4897
15.2925 | -105.4774
15.2962 | -105.4625
15.2999 | -105.447
15.3036 | -105.4307
15.3073 | -105.4144
15.311 | -105.3994
15.3147 | -105.3858
15.3184 | -105.3721
15.3221 | -105.3593
15.3257 | -105.3489
15.3294 | -105.3332
15.3331 | -105.3194
15.3368 | -105.3112
15.3405 | -105.3058
15.3442 | -105.3006
15.3479 | -105.2981
15.3516 | -105.297
15.3552 | -105.2965
15.3589 | -105.2969
15.3626 | -105.2995
15.3663 | -105.3066
15.37 | -105.3158
15.3737 | -105.3247
15.3774 | -105.3362
15.3811 | -105.3483
15.3848 | -105.3617
15.3884 | -105.3751
15.3921 | -105.3877
15.3958 | -105.4031
15.3995 | -105.4202
15.4032 | -105.4338
15.4069 | -105.4492
15.4106 | -105.4631
15.4143 | -105.4751
15.418 | -105.4879
15.4216 | -105.4993
15.4253 | -105.5088
15.429 | -105.5114
15.4327 | -105.5206
15.4364 | -105.5302
15.4401 | -105.5366
15.4438 | -105.5389
15.4475 | -105.5414
15.4512 | -105.5462
15.4548 | -105.552
15.4585 | -105.5526
15.4622 | -105.5498
15.4659 | -105.5452
15.4696 | -105.5418
15.4733 | -105.5386
15.477 | -105.534
15.4807 | -105.5229
15.4843 | -105.5095
15.488 | -105.4991
15.4917 | -105.4898
15.4954 | -105.4803
15.4991 | -105.4673
15.5028 | -105.4541
15.5065 | -105.4421
15.5102 | -105.4329
15.5139 | -105.4235
15.5175 | -105.4126
15.5212 | -105.4003
15.5249 | -105.3877
15.5286 | -105.3734
15.5323 | -105.358
15.536 | -105.3408
15.5397 | -105.3255
15.5434 | -105.3133
15.5471 | -105.2998
15.5507 | -105.2907
15.5544 | -105.2813
15.5581 | -105.2693
15.5618 | -105.259
15.5655 | -105.2506
15.5692 | -105.243
15.5729 | -105.2365
15.5766 | -105.2333
15.5802 | -105.2319
15.5839 | -105.2317
15.5876 | -105.2317
15.5913 | -105.2334
15.595 | -105.2361
15.5987 | -105.2355
15.6024 | -105.2378
15.6061 | -105.2428
15.6098 | -105.2497
15.6134 | -105.2568
15.6171 | -105.2651
15.6208 | -105.2705
15.6245 | -105.2776
15.6282 | -105.2834
15.6319 | -105.2891
15.6356 | -105.2973
15.6393 | -105.3036
15.643 | -105.3069
15.6466 | -105.3107
15.6503 | -105.3162
15.654 | -105.3166
15.6577 | -105.317
15.6614 | -105.3173
15.6651 | -105.3176
15.6688 | -105.3197
15.6725 | -105.3187
15.6761 | -105.3165
15.6798 | -105.3113
15.6835 | -105.3066
15.6872 | -105.3019
15.6909 | -105.2965
15.6946 | -105.2896
15.6983 | -105.2827
15.702 | -105.2772
15.7057 | -105.2732
15.7093 | -105.2669
15.713 | -105.2598
15.7167 | -105.2471
15.7204 | -105.2345
15.7241 | -105.2272
15.7278 | -105.2232
15.7315 | -105.2173
15.7352 | -105.2108
15.7389 | -105.2044
15.7425 | -105.1995
15.7462 | -105.1945
15.7499 | -105.1913
15.7536 | -105.1896
15.7573 | -105.1883
15.761 | -105.1898
15.7647 | -105.189
15.7684 | -105.185
15.7721 | -105.1812
15.7757 | -105.1784
15.7794 | -105.1745
15.7831 | -105.1722
15.7868 | -105.168
15.7905 | -105.1642
15.7942 | -105.1605
15.7979 | -105.1593
15.8016 | -105.155
15.8052 | -105.1545
15.8089 | -105.1534
15.8126 | -105.1516
15.8163 | -105.1498
15.82 | -105.1462
15.8237 | -105.1492
15.8274 | -105.1486
15.8311 | -105.1461
15.8348 | -105.1456
15.8384 | -105.1467
15.8421 | -105.1451
15.8458 | -105.1459
15.8495 | -105.1483
15.8532 | -105.1501
15.8569 | -105.151
15.8606 | -105.149
15.8643 | -105.15
15.868 | -105.1516
15.8716 | -105.1503
15.8753 | -105.1485
15.879 | -105.1475
15.8827 | -105.1487
15.8864 | -105.1482
15.8901 | -105.1474
15.8938 | -105.1456
15.8975 | -105.1457
15.9011 | -105.1453
15.9048 | -105.1485
15.9085 | -105.1517
15.9122 | -105.1525
15.9159 | -105.1517
15.9196 | -105.153
15.9233 | -105.1544
15.927 | -105.1551
15.9307 | -105.1557
15.9343 | -105.1603
15.938 | -105.1635
15.9417 | -105.1688
15.9454 | -105.1706
15.9491 | -105.1679
15.9528 | -105.161
15.9565 | -105.1584
15.9602 | -105.1554
15.9639 | -105.1535
15.9675 | -105.1529
15.9712 | -105.1518
15.9749 | -105.1516
15.9786 | -105.1478
15.9823 | -105.1426
15.986 | -105.1392
15.9897 | -105.1357
15.9934 | -105.1319
15.997 | -105.1296
16.0007 | -105.126
16.0044 | -105.123
16.0081 | -105.1221
16.0118 | -105.12
16.0155 | -105.1157
16.0192 | -105.1099
16.0229 | -105.1062
16.0266 | -105.0998
16.0302 | -105.0948
16.0339 | -105.0892
16.0376 | -105.0853
16.0413 | -105.0823
16.045 | -105.0795
16.0487 | -105.08
16.0524 | -105.0776
16.0561 | -105.0769
16.0598 | -105.0756
16.0634 | -105.0764
16.0671 | -105.0708
16.0708 | -105.0689
16.0745 | -105.071
16.0782 | -105.071
16.0819 | -105.0723
16.0856 | -105.0769
16.0893 | -105.084
16.0929 | -105.0901
16.0966 | -105.098
16.1003 | -105.1073
16.104 | -105.1135
16.1077 | -105.1221
16.1114 | -105.1282
16.1151 | -105.1357
16.1188 | -105.1464
16.1225 | -105.1582
16.1261 | -105.1681
16.1298 | -105.1783
16.1335 | -105.1881
16.1372 | -105.1982
16.1409 | -105.2104
16.1446 | -105.2233
16.1483 | -105.2342
16.152 | -105.2441
16.1557 | -105.2556
16.1593 | -105.267
16.163 | -105.2743
16.1667 | -105.2825
16.1704 | -105.2933
16.1741 | -105.3063
16.1778 | -105.3198
16.1815 | -105.3341
16.1852 | -105.3468
16.1889 | -105.3583
16.1925 | -105.3677
16.1962 | -105.3746
16.1999 | -105.3838
16.2036 | -105.3941
16.2073 | -105.404
16.211 | -105.4162
16.2147 | -105.4268
16.2184 | -105.434
16.222 | -105.4393
16.2257 | -105.4461
16.2294 | -105.4519
16.2331 | -105.4598
16.2368 | -105.4639
16.2405 | -105.4663
16.2442 | -105.4674
16.2479 | -105.4693
16.2516 | -105.473
16.2552 | -105.4738
16.2589 | -105.4726
16.2626 | -105.4697
16.2663 | -105.4646
16.27 | -105.4596
16.2737 | -105.4552
16.2774 | -105.4496
16.2811 | -105.4441
16.2848 | -105.4368
16.2884 | -105.426
16.2921 | -105.4181
16.2958 | -105.4104
16.2995 | -105.4076
16.3032 | -105.4012
16.3069 | -105.3921
16.3106 | -105.3841
16.3143 | -105.3764
16.3179 | -105.3703
16.3216 | -105.3645
16.3253 | -105.3585
16.329 | -105.3548
16.3327 | -105.3527
16.3364 | -105.3523
16.3401 | -105.3501
16.3438 | -105.3505
16.3475 | -105.3509
16.3511 | -105.3516
16.3548 | -105.3518
16.3585 | -105.3532
16.3622 | -105.3562
16.3659 | -105.3599
16.3696 | -105.3639
16.3733 | -105.3685
16.377 | -105.3751
16.3807 | -105.3835
16.3843 | -105.3935
16.388 | -105.4013
16.3917 | -105.4088
16.3954 | -105.4148
16.3991 | -105.4195
16.4028 | -105.4266
16.4065 | -105.4318
16.4102 | -105.4388
16.4139 | -105.4457
16.4175 | -105.449
16.4212 | -105.4538
16.4249 | -105.4611
16.4286 | -105.465
16.4323 | -105.4648
16.436 | -105.4649
16.4397 | -105.4664
16.4434 | -105.4659
16.447 | -105.4651
16.4507 | -105.4644
16.4544 | -105.4626
16.4581 | -105.4578
16.4618 | -105.4508
16.4655 | -105.4429
16.4692 | -105.4354
16.4729 | -105.4289
16.4766 | -105.4227
16.4802 | -105.4175
16.4839 | -105.4094
16.4876 | -105.3991
16.4913 | -105.391
16.495 | -105.3843
16.4987 | -105.3777
16.5024 | -105.3692
16.5061 | -105.3593
16.5098 | -105.3496
16.5134 | -105.3442
16.5171 | -105.3393
16.5208 | -105.3367
16.5245 | -105.3378
16.5282 | -105.3359
16.5319 | -105.3332
16.5356 | -105.333
16.5393 | -105.3329
16.5429 | -105.3337
16.5466 | -105.3375
16.5503 | -105.342
16.554 | -105.3448
16.5577 | -105.349
16.5614 | -105.3561
16.5651 | -105.3625
16.5688 | -105.3704
16.5725 | -105.3795
16.5761 | -105.3887
16.5798 | -105.399
16.5835 | -105.4101
16.5872 | -105.4189
16.5909 | -105.4312
16.5946 | -105.4454
16.5983 | -105.4582
16.602 | -105.4713
16.6057 | -105.4845
16.6093 | -105.4966
16.613 | -105.5064
16.6167 | -105.5141
16.6204 | -105.5253
16.6241 | -105.5385
16.6278 | -105.5508
16.6315 | -105.5668
16.6352 | -105.5833
16.6388 | -105.5948
16.6425 | -105.6066
16.6462 | -105.6167
16.6499 | -105.6268
16.6536 | -105.6377
16.6573 | -105.6484
16.661 | -105.6577
16.6647 | -105.6674
16.6684 | -105.6782
16.672 | -105.6896
16.6757 | -105.697
16.6794 | -105.7056
16.6831 | -105.7124
16.6868 | -105.7165
16.6905 | -105.7212
16.6942 | -105.7273
16.6979 | -105.7335
16.7016 | -105.7379
16.7052 | -105.7408
16.7089 | -105.7404
16.7126 | -105.7381
16.7163 | -105.7375
16.72 | -105.7337
16.7237 | -105.7339
16.7274 | -105.7319
16.7311 | -105.7287
16.7347 | -105.7226
16.7384 | -105.7185
16.7421 | -105.7166
16.7458 | -105.7117
16.7495 | -105.7045
16.7532 | -105.6956
16.7569 | -105.6899
16.7606 | -105.6819
16.7643 | -105.6736
16.7679 | -105.665
16.7716 | -105.6548
16.7753 | -105.6436
16.779 | -105.6348
16.7827 | -105.6271
16.7864 | -105.618
16.7901 | -105.6077
16.7938 | -105.5977
16.7975 | -105.5866
16.8011 | -105.5762
16.8048 | -105.5653
16.8085 | -105.5547
16.8122 | -105.5442
16.8159 | -105.5349
16.8196 | -105.5246
16.8233 | -105.515
16.827 | -105.5073
16.8307 | -105.4967
16.8343 | -105.4861
16.838 | -105.4768
16.8417 | -105.4673
16.8454 | -105.4586
16.8491 | -105.4487
16.8528 | -105.4418
16.8565 | -105.4381
16.8602 | -105.4315
16.8638 | -105.4237
16.8675 | -105.4181
16.8712 | -105.4112
16.8749 | -105.4047
16.8786 | -105.4013
16.8823 | -105.3984
16.886 | -105.3955
16.8897 | -105.3937
16.8934 | -105.3917
16.897 | -105.3903
16.9007 | -105.3879
16.9044 | -105.3858
16.9081 | -105.3852
16.9118 | -105.3844
16.9155 | -105.3847
16.9192 | -105.3839
16.9229 | -105.3857
16.9266 | -105.3865
16.9302 | -105.3879
16.9339 | -105.3887
16.9376 | -105.3885
16.9413 | -105.3896
16.945 | -105.3915
16.9487 | -105.3974
16.9524 | -105.4042
16.9561 | -105.4098
16.9597 | -105.4193
16.9634 | -105.4258
16.9671 | -105.4319
16.9708 | -105.4371
16.9745 | -105.4437
16.9782 | -105.4492
16.9819 | -105.4532
16.9856 | -105.4611
16.9893 | -105.4708
16.9929 | -105.4785
16.9966 | -105.4821
17.0003 | -105.4863
17.004 | -105.4908
17.0077 | -105.4945
17.0114 | -105.4965
17.0151 | -105.4977
17.0188 | -105.4944
17.0225 | -105.4907
17.0261 | -105.4815
17.0298 | -105.476
17.0335 | -105.47
17.0372 | -105.4587
17.0409 | -105.4478
17.0446 | -105.439
17.0483 | -105.4295
17.052 | -105.4204
17.0556 | -105.4123
17.0593 | -105.4031
17.063 | -105.3932
17.0667 | -105.3828
17.0704 | -105.3726
17.0741 | -105.3649
17.0778 | -105.3571
17.0815 | -105.3492
17.0852 | -105.343
17.0888 | -105.3377
17.0925 | -105.3342
17.0962 | -105.3302
17.0999 | -105.3247
17.1036 | -105.3187
17.1073 | -105.3152
17.111 | -105.3104
17.1147 | -105.306
17.1184 | -105.3035
17.122 | -105.3009
17.1257 | -105.2976
17.1294 | -105.2941
17.1331 | -105.2893
17.1368 | -105.2867
17.1405 | -105.2812
17.1442 | -105.279
17.1479 | -105.2792
17.1516 | -105.2774
17.1552 | -105.2772
17.1589 | -105.2778
17.1626 | -105.276
17.1663 | -105.2734
17.17 | -105.2746
17.1737 | -105.2724
17.1774 | -105.2697
17.1811 | -105.2676
17.1847 | -105.2645
17.1884 | -105.2627
17.1921 | -105.259
17.1958 | -105.2546
17.1995 | -105.2498
17.2032 | -105.2462
17.2069 | -105.2414
17.2106 | -105.2368
17.2143 | -105.2325
17.2179 | -105.2276
17.2216 | -105.2232
17.2253 | -105.2204
17.229 | -105.2154
17.2327 | -105.2085
17.2364 | -105.2025
17.2401 | -105.1971
17.2438 | -105.1925
17.2475 | -105.1885
17.2511 | -105.1864
17.2548 | -105.186
17.2585 | -105.1862
17.2622 | -105.1888
17.2659 | -105.189
17.2696 | -105.1907
17.2733 | -105.1933
17.277 | -105.1975
17.2806 | -105.2038
17.2843 | -105.2151
17.288 | -105.2266
17.2917 | -105.2355
17.2954 | -105.2455
17.2991 | -105.2563
17.3028 | -105.2691
17.3065 | -105.2832
17.3102 | -105.297
17.3138 | -105.3129
17.3175 | -105.3298
17.3212 | -105.3465
17.3249 | -105.3616
17.3286 | -105.3773
17.3323 | -105.3918
17.336 | -105.4066
17.3397 | -105.4231
17.3434 | -105.4383
17.347 | -105.452
17.3507 | -105.465
17.3544 | -105.4771
17.3581 | -105.4918
17.3618 | -105.5057
17.3655 | -105.5177
17.3692 | -105.5272
17.3729 | -105.5356
17.3766 | -105.5449
17.3802 | -105.5521
17.3839 | -105.5545
17.3876 | -105.5606
17.3913 | -105.5628
17.395 | -105.5626
17.3987 | -105.5592
17.4024 | -105.5526
17.4061 | -105.5478
17.4097 | -105.5464
17.4134 | -105.541
17.4171 | -105.5331
17.4208 | -105.5273
17.4245 | -105.5212
17.4282 | -105.512
17.4319 | -105.5004
17.4356 | -105.4862
17.4393 | -105.4703
17.4429 | -105.4543
17.4466 | -105.438
17.4503 | -105.4216
17.454 | -105.4081
17.4577 | -105.3937
17.4614 | -105.3777
17.4651 | -105.3626
17.4688 | -105.3447
17.4725 | -105.327
17.4761 | -105.3106
17.4798 | -105.2957
17.4835 | -105.2816
17.4872 | -105.2691
17.4909 | -105.2588
17.4946 | -105.2479
17.4983 | -105.236
17.502 | -105.2251
17.5056 | -105.214
17.5093 | -105.203
17.513 | -105.1943
17.5167 | -105.1857
17.5204 | -105.178
17.5241 | -105.1688
17.5278 | -105.1604
17.5315 | -105.154
17.5352 | -105.1447
17.5388 | -105.1374
17.5425 | -105.1332
17.5462 | -105.1274
17.5499 | -105.1186
17.5536 | -105.1126
17.5573 | -105.1081
17.561 | -105.1035
17.5647 | -105.0986
17.5684 | -105.0942
17.572 | -105.0911
17.5757 | -105.0902
17.5794 | -105.0893
17.5831 | -105.089
17.5868 | -105.0882
17.5905 | -105.0834
17.5942 | -105.0825
17.5979 | -105.0803
17.6015 | -105.0793
17.6052 | -105.078
17.6089 | -105.0743
17.6126 | -105.0725
17.6163 | -105.0719
17.62 | -105.0708
17.6237 | -105.0704
17.6274 | -105.0693
17.6311 | -105.0696
17.6347 | -105.0725
17.6384 | -105.0702
17.6421 | -105.0697
17.6458 | -105.0683
17.6495 | -105.0674
17.6532 | -105.0653
17.6569 | -105.0607
17.6606 | -105.0541
17.6643 | -105.0494
17.6679 | -105.0417
17.6716 | -105.0331
17.6753 | -105.0227
17.679 | -105.0159
17.6827 | -105.0083
17.6864 | -104.9993
17.6901 | -104.9907
17.6938 | -104.982
17.6974 | -104.9716
17.7011 | -104.9597
17.7048 | -104.9481
17.7085 | -104.9373
17.7122 | -104.9274
17.7159 | -104.917
17.7196 | -104.9065
17.7233 | -104.8951
17.727 | -104.8853
17.7306 | -104.8766
17.7343 | -104.8692
17.738 | -104.8621
17.7417 | -104.8549
17.7454 | -104.8507
17.7491 | -104.8491
17.7528 | -104.8466
17.7565 | -104.8477
17.7602 | -104.8491
17.7638 | -104.8529
17.7675 | -104.8561
17.7712 | -104.859
17.7749 | -104.8662
17.7786 | -104.8726
17.7823 | -104.8783
17.786 | -104.8868
17.7897 | -104.8939
17.7934 | -104.9014
17.797 | -104.9108
17.8007 | -104.923
17.8044 | -104.9342
17.8081 | -104.9439
17.8118 | -104.9541
17.8155 | -104.9648
17.8192 | -104.9747
17.8229 | -104.9815
17.8265 | -104.9889
17.8302 | -104.9947
17.8339 | -104.9995
17.8376 | -105.0057
17.8413 | -105.0116
17.845 | -105.0189
17.8487 | -105.0235
17.8524 | -105.0293
17.8561 | -105.0325
17.8597 | -105.0333
17.8634 | -105.0352
17.8671 | -105.0357
17.8708 | -105.0346
17.8745 | -105.0335
17.8782 | -105.0318
17.8819 | -105.0316
17.8856 | -105.0285
17.8893 | -105.024
17.8929 | -105.0212
17.8966 | -105.0181
17.9003 | -105.0149
17.904 | -105.0137
17.9077 | -105.0096
17.9114 | -105.0035
17.9151 | -104.998
17.9188 | -104.9926
17.9224 | -104.9893
17.9261 | -104.9822
17.9298 | -104.9736
17.9335 | -104.9594
17.9372 | -104.9463
17.9409 | -104.937
17.9446 | -104.9266
17.9483 | -104.9137
17.952 | -104.9016
17.9556 | -104.89
17.9593 | -104.8786
17.963 | -104.8664
17.9667 | -104.8525
17.9704 | -104.8399
17.9741 | -104.8279
17.9778 | -104.8159
17.9815 | -104.8038
17.9852 | -104.7895
17.9888 | -104.7726
17.9925 | -104.7569
17.9962 | -104.7398
17.9999 | -104.7228
18.0036 | -104.7054
18.0073 | -104.6891
18.011 | -104.6701
18.0147 | -104.6514
18.0184 | -104.6322
18.022 | -104.6148
18.0257 | -104.5979
18.0294 | -104.5807
18.0331 | -104.5635
18.0368 | -104.5458
18.0405 | -104.5281
18.0442 | -104.5125
18.0479 | -104.4953
18.0515 | -104.4777
18.0552 | -104.4607
18.0589 | -104.4413
18.0626 | -104.4201
18.0663 | -104.3997
18.07 | -104.3757
18.0737 | -104.3518
18.0774 | -104.328
18.0811 | -104.3035
18.0847 | -104.281
18.0884 | -104.2609
18.0921 | -104.2362
18.0958 | -104.2121
18.0995 | -104.1859
18.1032 | -104.1614
18.1069 | -104.1364
18.1106 | -104.1123
18.1143 | -104.0877
18.1179 | -104.0645
18.1216 | -104.0411
18.1253 | -104.0158
18.129 | -103.9921
18.1327 | -103.9689
18.1364 | -103.9459
18.1401 | -103.923
18.1438 | -103.9034
18.1474 | -103.8834
18.1511 | -103.863
18.1548 | -103.8431
18.1585 | -103.8252
18.1622 | -103.8076
18.1659 | -103.794
18.1696 | -103.7826
18.1733 | -103.7726
18.177 | -103.7628
18.1806 | -103.7543
18.1843 | -103.7495
18.188 | -103.7482
18.1917 | -103.7497
18.1954 | -103.7505
18.1991 | -103.7503
18.2028 | -103.7529
18.2065 | -103.7599
18.2102 | -103.7679
18.2138 | -103.7775
18.2175 | -103.7887
18.2212 | -103.8009
18.2249 | -103.8157
18.2286 | -103.8308
18.2323 | -103.8461
18.236 | -103.8633
18.2397 | -103.8797
18.2433 | -103.898
18.247 | -103.9203
18.2507 | -103.9414
18.2544 | -103.9645
18.2581 | -103.988
18.2618 | -104.0111
18.2655 | -104.0339
18.2692 | -104.0575
18.2729 | -104.0797
18.2765 | -104.1023
18.2802 | -104.1246
18.2839 | -104.1476
18.2876 | -104.1693
18.2913 | -104.1918
18.295 | -104.2134
18.2987 | -104.2361
18.3024 | -104.2577
18.3061 | -104.2803
18.3097 | -104.3033
18.3134 | -104.3238
18.3171 | -104.3448
18.3208 | -104.3683
18.3245 | -104.3892
18.3282 | -104.4116
18.3319 | -104.4331
18.3356 | -104.4568
18.3392 | -104.4787
18.3429 | -104.4963
18.3466 | -104.5106
18.3503 | -104.5249
18.354 | -104.5381
18.3577 | -104.5536
18.3614 | -104.5735
18.3651 | -104.5907
18.3688 | -104.604
18.3724 | -104.6154
18.3761 | -104.6271
18.3798 | -104.6363
18.3835 | -104.6436
18.3872 | -104.6513
18.3909 | -104.6572
18.3946 | -104.6643
18.3983 | -104.6706
18.402 | -104.6722
18.4056 | -104.6714
18.4093 | -104.6723
18.413 | -104.6741
18.4167 | -104.6743
18.4204 | -104.6715
18.4241 | -104.6683
18.4278 | -104.6657
18.4315 | -104.6616
18.4352 | -104.6541
18.4388 | -104.6465
18.4425 | -104.6369
18.4462 | -104.626
18.4499 | -104.6132
18.4536 | -104.6019
18.4573 | -104.5935
18.461 | -104.5845
18.4647 | -104.5735
18.4683 | -104.5606
18.472 | -104.5492
18.4757 | -104.538
18.4794 | -104.5275
18.4831 | -104.5197
18.4868 | -104.5134
18.4905 | -104.5059
18.4942 | -104.5003
18.4979 | -104.4952
18.5015 | -104.4912
18.5052 | -104.4875
18.5089 | -104.4833
18.5126 | -104.4812
18.5163 | -104.4798
18.52 | -104.4774
18.5237 | -104.4772
18.5274 | -104.479
18.5311 | -104.4829
18.5347 | -104.4884
18.5384 | -104.4913
18.5421 | -104.4944
18.5458 | -104.4976
18.5495 | -104.5024
18.5532 | -104.5089
18.5569 | -104.5175
18.5606 | -104.5256
18.5642 | -104.5333
18.5679 | -104.5419
18.5716 | -104.5502
18.5753 | -104.5601
18.579 | -104.5703
18.5827 | -104.5801
18.5864 | -104.5899
18.5901 | -104.5961
18.5938 | -104.6032
18.5974 | -104.6067
18.6011 | -104.609
18.6048 | -104.6144
18.6085 | -104.6208
18.6122 | -104.6283
18.6159 | -104.6374
18.6196 | -104.6438
18.6233 | -104.6497
18.627 | -104.6534
18.6306 | -104.6592
18.6343 | -104.6689
18.638 | -104.6772
18.6417 | -104.687
18.6454 | -104.6972
18.6491 | -104.706
18.6528 | -104.7147
18.6565 | -104.7231
18.6602 | -104.7298
18.6638 | -104.7376
18.6675 | -104.7461
18.6712 | -104.7534
18.6749 | -104.7596
18.6786 | -104.768
18.6823 | -104.7747
18.686 | -104.7793
18.6897 | -104.7862
18.6933 | -104.7912
18.697 | -104.7952
18.7007 | -104.7998
18.7044 | -104.8033
18.7081 | -104.8082
18.7118 | -104.8128
18.7155 | -104.8165
18.7192 | -104.8179
18.7229 | -104.8186
18.7265 | -104.8193
18.7302 | -104.8192
18.7339 | -104.817
18.7376 | -104.8142
18.7413 | -104.8113
18.745 | -104.8087
18.7487 | -104.8064
18.7524 | -104.8007
18.7561 | -104.7958
18.7597 | -104.7905
18.7634 | -104.7831
18.7671 | -104.7782
18.7708 | -104.7737
18.7745 | -104.7665
18.7782 | -104.7563
18.7819 | -104.7443
18.7856 | -104.736
18.7892 | -104.7251
18.7929 | -104.715
18.7966 | -104.7028
18.8003 | -104.69
18.804 | -104.6772
18.8077 | -104.6611
18.8114 | -104.644
18.8151 | -104.6271
18.8188 | -104.6075
18.8224 | -104.5891
18.8261 | -104.5686
18.8298 | -104.5462
18.8335 | -104.524
18.8372 | -104.5035
18.8409 | -104.4829
18.8446 | -104.4609
18.8483 | -104.4366
18.852 | -104.4136
18.8556 | -104.3921
18.8593 | -104.3674
18.863 | -104.343
18.8667 | -104.3201
18.8704 | -104.2948
18.8741 | -104.2691
18.8778 | -104.2425
18.8815 | -104.2169
18.8851 | -104.1907
18.8888 | -104.1645
18.8925 | -104.1389
18.8962 | -104.1123
18.8999 | -104.0863
18.9036 | -104.0623
18.9073 | -104.0416
18.911 | -104.0178
18.9147 | -103.9958
18.9183 | -103.9758
18.922 | -103.9575
18.9257 | -103.9402
18.9294 | -103.9251
18.9331 | -103.9097
18.9368 | -103.8971
18.9405 | -103.8874
18.9442 | -103.8781
18.9479 | -103.8681
18.9515 | -103.8613
18.9552 | -103.8579
18.9589 | -103.8511
18.9626 | -103.8457
18.9663 | -103.8421
18.97 | -103.8419
18.9737 | -103.8406
18.9774 | -103.8392
18.981 | -103.8392
18.9847 | -103.8416
18.9884 | -103.8445
18.9921 | -103.8465
18.9958 | -103.8485
18.9995 | -103.8528
19.0032 | -103.8595
19.0069 | -103.8656
19.0106 | -103.8703
19.0142 | -103.8725
19.0179 | -103.8765
19.0216 | -103.8837
19.0253 | -103.8892
19.029 | -103.8926
19.0327 | -103.8942
19.0364 | -103.8934
19.0401 | -103.8938
19.0438 | -103.8962
19.0474 | -103.8995
19.0511 | -103.9023
19.0548 | -103.9025
19.0585 | -103.9036
19.0622 | -103.9042
19.0659 | -103.9043
19.0696 | -103.9065
19.0733 | -103.9064
19.077 | -103.9062
19.0806 | -103.9053
19.0843 | -103.9006
19.088 | -103.8962
19.0917 | -103.8911
19.0954 | -103.8849
19.0991 | -103.8783
19.1028 | -103.8707
19.1065 | -103.8593
19.1101 | -103.8418
19.1138 | -103.8255
19.1175 | -103.8137
19.1212 | -103.8022
19.1249 | -103.7879
19.1286 | -103.7735
19.1323 | -103.7586
19.136 | -103.7428
19.1397 | -103.7245
19.1433 | -103.7041
19.147 | -103.6847
19.1507 | -103.6657
19.1544 | -103.6468
19.1581 | -103.6277
19.1618 | -103.6108
19.1655 | -103.595
19.1692 | -103.5787
19.1729 | -103.5594
19.1765 | -103.5383
19.1802 | -103.5197
19.1839 | -103.5053
19.1876 | -103.4912
19.1913 | -103.477
19.195 | -103.4611
19.1987 | -103.445
19.2024 | -103.4306
19.206 | -103.4169
19.2097 | -103.4056
19.2134 | -103.3938
19.2171 | -103.3832
19.2208 | -103.3738
19.2245 | -103.3657
19.2282 | -103.3578
19.2319 | -103.3552
19.2356 | -103.3519
19.2392 | -103.3486
19.2429 | -103.3452
19.2466 | -103.3435
19.2503 | -103.3433
19.254 | -103.3447
19.2577 | -103.3458
19.2614 | -103.3487
19.2651 | -103.3525
19.2688 | -103.3593
19.2724 | -103.3646
19.2761 | -103.3706
19.2798 | -103.3777
19.2835 | -103.3869
19.2872 | -103.3982
19.2909 | -103.408
19.2946 | -103.4174
19.2983 | -103.4297
19.302 | -103.4432
19.3056 | -103.4591
19.3093 | -103.4731
19.313 | -103.4856
19.3167 | -103.4972
19.3204 | -103.509
19.3241 | -103.5222
19.3278 | -103.5343
19.3315 | -103.5467
19.3351 | -103.5588
19.3388 | -103.5713
19.3425 | -103.5812
19.3462 | -103.5941
19.3499 | -103.6095
19.3536 | -103.6233
19.3573 | -103.6372
19.361 | -103.6506
19.3647 | -103.6642
19.3683 | -103.6781
19.372 | -103.6914
19.3757 | -103.7054
19.3794 | -103.7213
19.3831 | -103.7414
19.3868 | -103.7575
19.3905 | -103.7685
19.3942 | -103.7788
19.3979 | -103.7905
19.4015 | -103.8038
19.4052 | -103.8179
19.4089 | -103.8301
19.4126 | -103.8445
19.4163 | -103.8605
19.42 | -103.8757
19.4237 | -103.8904
19.4274 | -103.9019
19.431 | -103.9121
19.4347 | -103.9252
19.4384 | -103.9366
19.4421 | -103.9485
19.4458 | -103.9615
19.4495 | -103.9776
19.4532 | -103.994
19.4569 | -104.007
19.4606 | -104.0193
19.4642 | -104.0302
19.4679 | -104.0426
19.4716 | -104.0547
19.4753 | -104.07
19.479 | -104.0836
19.4827 | -104.0969
19.4864 | -104.1084
19.4901 | -104.1215
19.4938 | -104.1331
19.4974 | -104.1449
19.5011 | -104.1576
19.5048 | -104.1712
19.5085 | -104.182
19.5122 | -104.1944
19.5159 | -104.208
19.5196 | -104.22
19.5233 | -104.231
19.5269 | -104.241
19.5306 | -104.2481
19.5343 | -104.2543
19.538 | -104.2608
19.5417 | -104.268
19.5454 | -104.2722
19.5491 | -104.2777
19.5528 | -104.281
19.5565 | -104.2807
19.5601 | -104.2815
19.5638 | -104.2831
19.5675 | -104.284
19.5712 | -104.2823
19.5749 | -104.2798
19.5786 | -104.2784
19.5823 | -104.2745
19.586 | -104.2732
19.5897 | -104.2707
19.5933 | -104.2673
19.597 | -104.2647
19.6007 | -104.2599
19.6044 | -104.2561
19.6081 | -104.2523
19.6118 | -104.2473
19.6155 | -104.2404
19.6192 | -104.2366
19.6228 | -104.2339
19.6265 | -104.2278
19.6302 | -104.2227
19.6339 | -104.2194
19.6376 | -104.2159
19.6413 | -104.2098
19.645 | -104.2041
19.6487 | -104.1996
19.6524 | -104.1937
19.656 | -104.1873
19.6597 | -104.1807
19.6634 | -104.1763
19.6671 | -104.1727
19.6708 | -104.1663
19.6745 | -104.16
19.6782 | -104.1537
19.6819 | -104.1486
19.6856 | -104.1449
19.6892 | -104.1392
19.6929 | -104.1337
19.6966 | -104.1288
19.7003 | -104.1219
19.704 | -104.1148
19.7077 | -104.1109
19.7114 | -104.1073
19.7151 | -104.1027
19.7188 | -104.0984
19.7224 | -104.092
19.7261 | -104.0855
19.7298 | -104.0785
19.7335 | -104.0736
19.7372 | -104.0689
19.7409 | -104.0633
19.7446 | -104.0597
19.7483 | -104.0576
19.7519 | -104.0581
19.7556 | -104.0548
19.7593 | -104.0511
19.763 | -104.049
19.7667 | -104.0465
19.7704 | -104.0434
19.7741 | -104.0418
19.7778 | -104.0394
19.7815 | -104.0376
19.7851 | -104.0373
19.7888 | -104.037
19.7925 | -104.037
19.7962 | -104.0368
19.7999 | -104.0372
19.8036 | -104.0404
19.8073 | -104.0429
19.811 | -104.0456
19.8147 | -104.0483
19.8183 | -104.0516
19.822 | -104.054
19.8257 | -104.0556
19.8294 | -104.058
19.8331 | -104.0629
19.8368 | -104.0687
19.8405 | -104.0715
19.8442 | -104.0738
19.8478 | -104.0756
19.8515 | -104.0784
19.8552 | -104.0814
19.8589 | -104.0844
19.8626 | -104.0878
19.8663 | -104.0875
19.87 | -104.0878
19.8737 | -104.0896
19.8774 | -104.091
19.881 | -104.0929
19.8847 | -104.0963
19.8884 | -104.0969
19.8921 | -104.0974
19.8958 | -104.0985
19.8995 | -104.0976
19.9032 | -104.0963
19.9069 | -104.0968
19.9106 | -104.0949
19.9142 | -104.0914
19.9179 | -104.0905
19.9216 | -104.0897
19.9253 | -104.0887
19.929 | -104.0861
19.9327 | -104.0823
19.9364 | -104.0807
19.9401 | -104.0794
19.9438 | -104.0772
19.9474 | -104.0712
19.9511 | -104.0671
19.9548 | -104.0631
19.9585 | -104.0608
19.9622 | -104.0565
19.9659 | -104.0526
19.9696 | -104.0495
19.9733 | -104.0491
19.9769 | -104.0495
19.9806 | -104.0491
19.9843 | -104.0493
19.988 | -104.0484
19.9917 | -104.0479
19.9954 | -104.0505
19.9991 | -104.0511
20.0028 | -104.0517
20.0065 | -104.0528
20.0101 | -104.0544
20.0138 | -104.0559
20.0175 | -104.0556
20.0212 | -104.0572
20.0249 | -104.058
20.0286 | -104.0602
20.0323 | -104.059
20.036 | -104.0576
20.0397 | -104.0585
20.0433 | -104.0593
20.047 | -104.06
20.0507 | -104.0612
20.0544 | -104.061
20.0581 | -104.0594
20.0618 | -104.0589
20.0655 | -104.0554
20.0692 | -104.0489
20.0728 | -104.0438
20.0765 | -104.0372
20.0802 | -104.0294
20.0839 | -104.021
20.0876 | -104.0128
20.0913 | -104.0056
20.095 | -103.9985
20.0987 | -103.9893
20.1024 | -103.9782
20.106 | -103.9675
20.1097 | -103.9567
20.1134 | -103.9456
20.1171 | -103.9332
20.1208 | -103.9221
20.1245 | -103.912
20.1282 | -103.8977
20.1319 | -103.8846
20.1356 | -103.8742
20.1392 | -103.8639
20.1429 | -103.8522
20.1466 | -103.8415
20.1503 | -103.833
20.154 | -103.8251
20.1577 | -103.8164
20.1614 | -103.8048
20.1651 | -103.7944
20.1687 | -103.7838
20.1724 | -103.7749
20.1761 | -103.7668
20.1798 | -103.7572
20.1835 | -103.7494
20.1872 | -103.7408
20.1909 | -103.7336
20.1946 | -103.7283
20.1983 | -103.7235
20.2019 | -103.7178
20.2056 | -103.7125
20.2093 | -103.7106
20.213 | -103.7064
20.2167 | -103.6991
20.2204 | -103.6949
20.2241 | -103.6889
20.2278 | -103.6861
20.2315 | -103.6835
20.2351 | -103.6797
20.2388 | -103.6752
20.2425 | -103.6715
20.2462 | -103.6668
20.2499 | -103.6625
20.2536 | -103.6598
20.2573 | -103.6576
20.261 | -103.652
20.2646 | -103.6473
20.2683 | -103.6442
20.272 | -103.6406
20.2757 | -103.6392
20.2794 | -103.6362
20.2831 | -103.6311
20.2868 | -103.6278
20.2905 | -103.624
20.2942 | -103.6197
20.2978 | -103.6168
20.3015 | -103.6108
20.3052 | -103.6061
20.3089 | -103.6009
20.3126 | -103.5975
20.3163 | -103.5916
20.32 | -103.5865
20.3237 | -103.5822
20.3274 | -103.5761
20.331 | -103.5709
20.3347 | -103.5653
20.3384 | -103.561
20.3421 | -103.5569
20.3458 | -103.5535
20.3495 | -103.5483
20.3532 | -103.5453
20.3569 | -103.5444
20.3606 | -103.5426
20.3642 | -103.5423
20.3679 | -103.5445
20.3716 | -103.547
20.3753 | -103.5468
20.379 | -103.5452
20.3827 | -103.5467
20.3864 | -103.5515
20.3901 | -103.5545
20.3937 | -103.5558
20.3974 | -103.5595
20.4011 | -103.5602
20.4048 | -103.5609
20.4085 | -103.5594
20.4122 | -103.5565
20.4159 | -103.5553
20.4196 | -103.5559
20.4233 | -103.5538
20.4269 | -103.5493
20.4306 | -103.5436
20.4343 | -103.5361
20.438 | -103.5281
20.4417 | -103.5179
20.4454 | -103.509
20.4491 | -103.501
20.4528 | -103.486
20.4565 | -103.4682
20.4601 | -103.4515
20.4638 | -103.4345
20.4675 | -103.4166
20.4712 | -103.3967
20.4749 | -103.3728
20.4786 | -103.3466
20.4823 | -103.3176
20.486 | -103.2865
20.4896 | -103.2541
20.4933 | -103.2213
20.497 | -103.1848
20.5007 | -103.1511
20.5044 | -103.1186
20.5081 | -103.08
20.5118 | -103.0392
20.5155 | -102.9984
20.5192 | -102.9542
20.5228 | -102.9081
20.5265 | -102.8663
20.5302 | -102.8243
20.5339 | -102.7782
20.5376 | -102.7322
20.5413 | -102.6859
20.545 | -102.6372
20.5487 | -102.5897
20.5524 | -102.5443
20.556 | -102.4981
20.5597 | -102.4526
20.5634 | -102.4052
20.5671 | -102.3602
20.5708 | -102.3183
20.5745 | -102.2743
20.5782 | -102.2316
20.5819 | -102.1908
20.5856 | -102.1499
20.5892 | -102.1095
20.5929 | -102.0683
20.5966 | -102.0234
20.6003 | -101.9819
20.604 | -101.9386
20.6077 | -101.8934
20.6114 | -101.8511
20.6151 | -101.8094
20.6187 | -101.7683
20.6224 | -101.7244
20.6261 | -101.6779
20.6298 | -101.6309
20.6335 | -101.5851
20.6372 | -101.5426
20.6409 | -101.5033
20.6446 | -101.4609
20.6483 | -101.4156
20.6519 | -101.3719
20.6556 | -101.3298
20.6593 | -101.2889
20.663 | -101.2439
20.6667 | -101.2006
20.6704 | -101.1603
20.6741 | -101.1177
20.6778 | -101.0731
20.6815 | -101.0312
20.6851 | -100.991
20.6888 | -100.9515
20.6925 | -100.9119
20.6962 | -100.8736
20.6999 | -100.8372
20.7036 | -100.8013
20.7073 | -100.7665
20.711 | -100.7282
20.7146 | -100.6885
20.7183 | -100.6527
20.722 | -100.6181
20.7257 | -100.5871
20.7294 | -100.5592
20.7331 | -100.5327
20.7368 | -100.5064
20.7405 | -100.4813
20.7442 | -100.461
20.7478 | -100.4426
20.7515 | -100.4339
20.7552 | -100.4158
20.7589 | -100.3988
20.7626 | -100.3848
20.7663 | -100.3721
20.77 | -100.363
20.7737 | -100.3531
20.7774 | -100.3485
20.781 | -100.348
20.7847 | -100.3478
20.7884 | -100.3456
20.7921 | -100.3439
20.7958 | -100.3469
20.7995 | -100.3513
20.8032 | -100.3569
20.8069 | -100.3647
20.8105 | -100.3724
20.8142 | -100.3806
20.8179 | -100.3917
20.8216 | -100.4019
20.8253 | -100.4123
20.829 | -100.4251
20.8327 | -100.4358
20.8364 | -100.4485
20.8401 | -100.4614
20.8437 | -100.4725
20.8474 | -100.4851
20.8511 | -100.4973
20.8548 | -100.5088
20.8585 | -100.5219
20.8622 | -100.533
20.8659 | -100.5428
20.8696 | -100.5541
20.8733 | -100.5654
20.8769 | -100.5752
20.8806 | -100.5844
20.8843 | -100.5926
20.888 | -100.6022
20.8917 | -100.6124
20.8954 | -100.6189
20.8991 | -100.627
20.9028 | -100.6369
20.9064 | -100.6451
20.9101 | -100.6532
20.9138 | -100.66
20.9175 | -100.6698
20.9212 | -100.6799
20.9249 | -100.6891
20.9286 | -100.6969
20.9323 | -100.7036
20.936 | -100.7089
20.9396 | -100.7171
20.9433 | -100.7271
20.947 | -100.7356
20.9507 | -100.7419
20.9544 | -100.7471
20.9581 | -100.7558
20.9618 | -100.764
20.9655 | -100.771
20.9692 | -100.7791
20.9728 | -100.7898
20.9765 | -100.7997
20.9802 | -100.8076
20.9839 | -100.8148
20.9876 | -100.8232
20.9913 | -100.8304
20.995 | -100.8371
20.9987 | -100.8441
21.0024 | -100.8517
21.006 | -100.862
21.0097 | -100.8736
21.0134 | -100.8813
21.0171 | -100.8882
21.0208 | -100.8938
21.0245 | -100.8983
21.0282 | -100.902
21.0319 | -100.9052
21.0355 | -100.9105
21.0392 | -100.9139
21.0429 | -100.9165
21.0466 | -100.9232
21.0503 | -100.9322
21.054 | -100.9382
21.0577 | -100.9442
21.0614 | -100.9509
21.0651 | -100.9567
21.0687 | -100.9623
21.0724 | -100.9717
21.0761 | -100.9782
21.0798 | -100.9832
21.0835 | -100.9901
21.0872 | -101.0011
21.0909 | -101.0136
21.0946 | -101.0204
21.0983 | -101.0286
21.1019 | -101.0381
21.1056 | -101.0474
21.1093 | -101.0599
21.113 | -101.0732
21.1167 | -101.0857
21.1204 | -101.0968
21.1241 | -101.1083
21.1278 | -101.1209
21.1314 | -101.1338
21.1351 | -101.1451
21.1388 | -101.159
21.1425 | -101.1727
21.1462 | -101.1832
21.1499 | -101.1935
21.1536 | -101.2055
21.1573 | -101.2195
21.161 | -101.2327
21.1646 | -101.2463
21.1683 | -101.2596
21.172 | -101.2717
21.1757 | -101.2854
21.1794 | -101.299
21.1831 | -101.3135
21.1868 | -101.3285
21.1905 | -101.3424
21.1942 | -101.3576
21.1978 | -101.3698
21.2015 | -101.3799
21.2052 | -101.389
21.2089 | -101.3981
21.2126 | -101.4082
21.2163 | -101.4173
21.22 | -101.4274
21.2237 | -101.4362
21.2274 | -101.4444
21.231 | -101.4506
21.2347 | -101.4554
21.2384 | -101.46
21.2421 | -101.4636
21.2458 | -101.4677
21.2495 | -101.4714
21.2532 | -101.4741
21.2569 | -101.477
21.2605 | -101.4785
21.2642 | -101.4778
21.2679 | -101.478
21.2716 | -101.4771
21.2753 | -101.477
21.279 | -101.4766
21.2827 | -101.4751
21.2864 | -101.4756
21.2901 | -101.4753
21.2937 | -101.4758
21.2974 | -101.4766
21.3011 | -101.4776
21.3048 | -101.4775
21.3085 | -101.4772
21.3122 | -101.4774
21.3159 | -101.4792
21.3196 | -101.4825
21.3233 | -101.4834
21.3269 | -101.4842
21.3306 | -101.4893
21.3343 | -101.4939
21.338 | -101.4986
21.3417 | -101.5051
21.3454 | -101.5128
21.3491 | -101.518
21.3528 | -101.5226
21.3564 | -101.528
21.3601 | -101.5329
21.3638 | -101.5361
21.3675 | -101.5412
21.3712 | -101.5465
21.3749 | -101.551
21.3786 | -101.5562
21.3823 | -101.5614
21.386 | -101.5657
21.3896 | -101.5691
21.3933 | -101.5712
21.397 | -101.5745
21.4007 | -101.5777
21.4044 | -101.5815
21.4081 | -101.5838
21.4118 | -101.5832
21.4155 | -101.5818
21.4192 | -101.5781
21.4228 | -101.5757
21.4265 | -101.5738
21.4302 | -101.5687
21.4339 | -101.5628
21.4376 | -101.5591
21.4413 | -101.551
21.445 | -101.5408
21.4487 | -101.5343
21.4523 | -101.5282
21.456 | -101.5167
21.4597 | -101.5015
21.4634 | -101.4851
21.4671 | -101.4678
21.4708 | -101.4516
21.4745 | -101.4341
21.4782 | -101.4146
21.4819 | -101.3958
21.4855 | -101.3765
21.4892 | -101.3559
21.4929 | -101.3375
21.4966 | -101.3174
21.5003 | -101.2972
21.504 | -101.2775
21.5077 | -101.2583
21.5114 | -101.2358
21.5151 | -101.2144
21.5187 | -101.1932
21.5224 | -101.1712
21.5261 | -101.1492
21.5298 | -101.1261
21.5335 | -101.1042
21.5372 | -101.0838
21.5409 | -101.0629
21.5446 | -101.0423
21.5482 | -101.0209
21.5519 | -101.0006
21.5556 | -100.9828
21.5593 | -100.9651
21.563 | -100.948
21.5667 | -100.9322
21.5704 | -100.916
21.5741 | -100.9001
21.5778 | -100.8813
21.5814 | -100.8671
21.5851 | -100.8521
21.5888 | -100.84
21.5925 | -100.8257
21.5962 | -100.8141
21.5999 | -100.8042
21.6036 | -100.7955
21.6073 | -100.7879
21.611 | -100.7816
21.6146 | -100.7752
21.6183 | -100.7709
21.622 | -100.7689
21.6257 | -100.7678
21.6294 | -100.7675
21.6331 | -100.7684
21.6368 | -100.7691
21.6405 | -100.7736
21.6442 | -100.7791
21.6478 | -100.7848
21.6515 | -100.7904
21.6552 | -100.7958
21.6589 | -100.8027
21.6626 | -100.813
21.6663 | -100.8228
21.67 | -100.8343
21.6737 | -100.8468
21.6773 | -100.858
21.681 | -100.8695
21.6847 | -100.882
21.6884 | -100.8947
21.6921 | -100.9084
21.6958 | -100.9239
21.6995 | -100.9389
21.7032 | -100.9551
21.7069 | -100.9711
21.7105 | -100.9866
21.7142 | -101.0016
21.7179 | -101.0157
21.7216 | -101.032
21.7253 | -101.0506
21.729 | -101.0687
21.7327 | -101.0846
21.7364 | -101.1001
21.7401 | -101.1152
21.7437 | -101.1267
21.7474 | -101.1403
21.7511 | -101.1549
21.7548 | -101.1671
21.7585 | -101.1804
21.7622 | -101.1959
21.7659 | -101.2103
21.7696 | -101.2252
21.7732 | -101.2385
21.7769 | -101.2479
21.7806 | -101.2566
21.7843 | -101.2667
21.788 | -101.2773
21.7917 | -101.2858
21.7954 | -101.2922
21.7991 | -101.2974
21.8028 | -101.3011
21.8064 | -101.3047
21.8101 | -101.3085
21.8138 | -101.3105
21.8175 | -101.3105
21.8212 | -101.3137
21.8249 | -101.3171
21.8286 | -101.3183
21.8323 | -101.3189
21.836 | -101.3206
21.8396 | -101.3241
21.8433 | -101.3251
21.847 | -101.3252
21.8507 | -101.3259
21.8544 | -101.3277
21.8581 | -101.3281
21.8618 | -101.3285
21.8655 | -101.3273
21.8692 | -101.3257
21.8728 | -101.3265
21.8765 | -101.3284
21.8802 | -101.329
21.8839 | -101.3317
21.8876 | -101.3387
21.8913 | -101.3432
21.895 | -101.3492
21.8987 | -101.3537
21.9023 | -101.3594
21.906 | -101.3658
21.9097 | -101.3728
21.9134 | -101.3808
21.9171 | -101.3913
21.9208 | -101.4032
21.9245 | -101.4172
21.9282 | -101.4309
21.9319 | -101.441
21.9355 | -101.452
21.9392 | -101.4667
21.9429 | -101.4822
21.9466 | -101.4973
21.9503 | -101.5103
21.954 | -101.5246
21.9577 | -101.5424
21.9614 | -101.5615
21.9651 | -101.5829
21.9687 | -101.6042
21.9724 | -101.6234
21.9761 | -101.6449
21.9798 | -101.6628
21.9835 | -101.6756
21.9872 | -101.6926
21.9909 | -101.7113
21.9946 | -101.7292
21.9982 | -101.7462
22.0019 | -101.7609
22.0056 | -101.7747
22.0093 | -101.7915
22.013 | -101.8063
22.0167 | -101.8198
22.0204 | -101.8342
22.0241 | -101.8473
22.0278 | -101.8596
22.0314 | -101.8713
22.0351 | -101.8804
22.0388 | -101.8894
22.0425 | -101.8983
22.0462 | -101.9083
22.0499 | -101.9181
22.0536 | -101.9289
22.0573 | -101.9357
22.061 | -101.9417
22.0646 | -101.9463
22.0683 | -101.9498
22.072 | -101.9517
22.0757 | -101.9556
22.0794 | -101.9606
22.0831 | -101.9648
22.0868 | -101.9703
22.0905 | -101.9767
22.0941 | -101.9802
22.0978 | -101.983
22.1015 | -101.9878
22.1052 | -101.9946
22.1089 | -102.0007
22.1126 | -102.0087
22.1163 | -102.0169
22.12 | -102.023
22.1237 | -102.0286
22.1273 | -102.0353
22.131 | -102.0439
22.1347 | -102.053
22.1384 | -102.0624
22.1421 | -102.0737
22.1458 | -102.0857
22.1495 | -102.0964
22.1532 | -102.1065
22.1569 | -102.1187
22.1605 | -102.1325
22.1642 | -102.1451
22.1679 | -102.1583
22.1716 | -102.175
22.1753 | -102.1912
22.179 | -102.206
22.1827 | -102.2206
22.1864 | -102.2335
22.19 | -102.2461
22.1937 | -102.2608
22.1974 | -102.2761
22.2011 | -102.2897
22.2048 | -102.3046
22.2085 | -102.32
22.2122 | -102.3364
22.2159 | -102.3556
22.2196 | -102.3734
22.2232 | -102.3895
22.2269 | -102.4056
22.2306 | -102.4211
22.2343 | -102.4369
22.238 | -102.4534
22.2417 | -102.4712
22.2454 | -102.4921
22.2491 | -102.5133
22.2528 | -102.5324
22.2564 | -102.5517
22.2601 | -102.5718
22.2638 | -102.5943
22.2675 | -102.6135
22.2712 | -102.6308
22.2749 | -102.6495
22.2786 | -102.6677
22.2823 | -102.6851
22.286 | -102.7045
22.2896 | -102.725
22.2933 | -102.748
22.297 | -102.776
22.3007 | -102.8012
22.3044 | -102.8226
22.3081 | -102.8431
22.3118 | -102.8647
22.3155 | -102.8889
22.3191 | -102.9144
22.3228 | -102.9354
22.3265 | -102.9556
22.3302 | -102.9775
22.3339 | -102.9993
22.3376 | -103.0223
22.3413 | -103.0459
22.345 | -103.0675
22.3487 | -103.087
22.3523 | -103.108
22.356 | -103.1294
22.3597 | -103.1505
22.3634 | -103.1709
22.3671 | -103.1926
22.3708 | -103.2123
22.3745 | -103.2348
22.3782 | -103.2572
22.3819 | -103.2767
22.3855 | -103.2938
22.3892 | -103.3128
22.3929 | -103.3298
22.3966 | -103.3418
22.4003 | -103.356
22.404 | -103.3708
22.4077 | -103.3832
22.4114 | -103.3951
22.415 | -103.4056
22.4187 | -103.415
22.4224 | -103.4226
22.4261 | -103.4306
22.4298 | -103.4345
22.4335 | -103.44
22.4372 | -103.4461
22.4409 | -103.4486
22.4446 | -103.4453
22.4482 | -103.4388
22.4519 | -103.4294
22.4556 | -103.4136
22.4593 | -103.3997
22.463 | -103.3872
22.4667 | -103.3748
22.4704 | -103.3625
22.4741 | -103.3464
22.4778 | -103.3279
22.4814 | -103.3093
22.4851 | -103.2888
22.4888 | -103.2665
22.4925 | -103.2439
22.4962 | -103.2211
22.4999 | -103.1975
22.5036 | -103.1703
22.5073 | -103.1416
22.511 | -103.1121
22.5146 | -103.078
22.5183 | -103.0454
22.522 | -103.0137
22.5257 | -102.98
22.5294 | -102.9429
22.5331 | -102.9034
22.5368 | -102.8631
22.5405 | -102.8252
22.5441 | -102.7868
22.5478 | -102.7494
22.5515 | -102.7128
22.5552 | -102.6768
22.5589 | -102.6404
22.5626 | -102.6053
22.5663 | -102.5709
22.57 | -102.5369
22.5737 | -102.504
22.5773 | -102.4734
22.581 | -102.4451
22.5847 | -102.4184
22.5884 | -102.3935
22.5921 | -102.3705
22.5958 | -102.3498
22.5995 | -102.3306
22.6032 | -102.3127
22.6069 | -102.2965
22.6105 | -102.2831
22.6142 | -102.2718
22.6179 | -102.2608
22.6216 | -102.2534
22.6253 | -102.2493
22.629 | -102.2479
22.6327 | -102.2469
22.6364 | -102.2482
22.64 | -102.2525
22.6437 | -102.2604
22.6474 | -102.2693
22.6511 | -102.2782
22.6548 | -102.2903
22.6585 | -102.3048
22.6622 | -102.3211
22.6659 | -102.34
22.6696 | -102.3599
22.6732 | -102.381
22.6769 | -102.4012
22.6806 | -102.4225
22.6843 | -102.4456
22.688 | -102.4691
22.6917 | -102.4924
22.6954 | -102.5136
22.6991 | -102.5356
22.7028 | -102.5591
22.7064 | -102.5839
22.7101 | -102.6087
22.7138 | -102.6326
22.7175 | -102.659
22.7212 | -102.6853
22.7249 | -102.7089
22.7286 | -102.7324
22.7323 | -102.7556
22.7359 | -102.7765
22.7396 | -102.7973
22.7433 | -102.817
22.747 | -102.8354
22.7507 | -102.8553
22.7544 | -102.8723
22.7581 | -102.8888
22.7618 | -102.9049
22.7655 | -102.9221
22.7691 | -102.9385
22.7728 | -102.9565
22.7765 | -102.9742
22.7802 | -102.9852
22.7839 | -102.9941
22.7876 | -103.0006
22.7913 | -103.0059
22.795 | -103.0128
22.7987 | -103.0163
22.8023 | -103.0173
22.806 | -103.0174
22.8097 | -103.0141
22.8134 | -103.0101
22.8171 | -103.006
22.8208 | -103.0028
22.8245 | -102.9984
22.8282 | -102.9932
22.8318 | -102.9882
22.8355 | -102.9827
22.8392 | -102.9743
22.8429 | -102.965
22.8466 | -102.9557
22.8503 | -102.9462
22.854 | -102.9376
22.8577 | -102.9286
22.8614 | -102.9159
22.865 | -102.9016
22.8687 | -102.8873
22.8724 | -102.8701
22.8761 | -102.8527
22.8798 | -102.8327
22.8835 | -102.8157
22.8872 | -102.7995
22.8909 | -102.7831
22.8946 | -102.7658
22.8982 | -102.7472
22.9019 | -102.7284
22.9056 | -102.71
22.9093 | -102.6919
22.913 | -102.6729
22.9167 | -102.6547
22.9204 | -102.6349
22.9241 | -102.614
22.9278 | -102.5955
22.9314 | -102.5783
22.9351 | -102.5611
22.9388 | -102.5442
22.9425 | -102.53
22.9462 | -102.5139
22.9499 | -102.496
22.9536 | -102.4784
22.9573 | -102.4612
22.9609 | -102.4447
22.9646 | -102.429
22.9683 | -102.4125
22.972 | -102.3982
22.9757 | -102.3855
22.9794 | -102.3723
22.9831 | -102.36
22.9868 | -102.3478
22.9905 | -102.3357
22.9941 | -102.3236
22.9978 | -102.3123
23.0015 | -102.3034
23.0052 | -102.2936
23.0089 | -102.2831
23.0126 | -102.2731
23.0163 | -102.267
23.02 | -102.2636
23.0237 | -102.2593
23.0273 | -102.2558
23.031 | -102.2526
23.0347 | -102.2502
23.0384 | -102.2501
23.0421 | -102.2483
23.0458 | -102.246
23.0495 | -102.2457
23.0532 | -102.246
23.0568 | -102.2469
23.0605 | -102.25
23.0642 | -102.2513
23.0679 | -102.2547
23.0716 | -102.2587
23.0753 | -102.2615
23.079 | -102.2655
23.0827 | -102.2699
23.0864 | -102.274
23.09 | -102.2769
23.0937 | -102.2812
23.0974 | -102.2855
23.1011 | -102.2905
23.1048 | -102.2975
23.1085 | -102.3033
23.1122 | -102.3087
23.1159 | -102.3136
23.1196 | -102.3155
23.1232 | -102.3187
23.1269 | -102.3252
23.1306 | -102.3299
23.1343 | -102.3359
23.138 | -102.3416
23.1417 | -102.3481
23.1454 | -102.3531
23.1491 | -102.3589
23.1528 | -102.3636
23.1564 | -102.3672
23.1601 | -102.3707
23.1638 | -102.3756
23.1675 | -102.3799
23.1712 | -102.3859
23.1749 | -102.3924
23.1786 | -102.3972
23.1823 | -102.402
23.1859 | -102.4098
23.1896 | -102.417
23.1933 | -102.4224
23.197 | -102.4296
23.2007 | -102.4372
23.2044 | -102.4435
23.2081 | -102.4506
23.2118 | -102.4594
23.2155 | -102.4654
23.2191 | -102.4718
23.2228 | -102.4801
23.2265 | -102.488
23.2302 | -102.4945
23.2339 | -102.5025
23.2376 | -102.5094
23.2413 | -102.5169
23.245 | -102.5243
23.2487 | -102.534
23.2523 | -102.5449
23.256 | -102.5528
23.2597 | -102.5576
23.2634 | -102.5625
23.2671 | -102.5675
23.2708 | -102.5719
23.2745 | -102.5768
23.2782 | -102.5843
23.2818 | -102.5916
23.2855 | -102.5989
23.2892 | -102.6039
23.2929 | -102.6087
23.2966 | -102.6127
23.3003 | -102.6142
23.304 | -102.6167
23.3077 | -102.6193
23.3114 | -102.6201
23.315 | -102.6204
23.3187 | -102.6217
23.3224 | -102.6227
23.3261 | -102.6243
23.3298 | -102.6251
23.3335 | -102.625
23.3372 | -102.6242
23.3409 | -102.6218
23.3446 | -102.6176
23.3482 | -102.6133
23.3519 | -102.608
23.3556 | -102.6027
23.3593 | -102.5964
23.363 | -102.5896
23.3667 | -102.5831
23.3704 | -102.5784
23.3741 | -102.5724
23.3777 | -102.5682
23.3814 | -102.5644
23.3851 | -102.5591
23.3888 | -102.5517
23.3925 | -102.5466
23.3962 | -102.5418
23.3999 | -102.537
23.4036 | -102.5337
23.4073 | -102.5279
23.4109 | -102.5202
23.4146 | -102.5128
23.4183 | -102.5078
23.422 | -102.505
23.4257 | -102.4994
23.4294 | -102.4929
23.4331 | -102.4878
23.4368 | -102.4837
23.4405 | -102.4783
23.4441 | -102.4726
23.4478 | -102.4677
23.4515 | -102.4639
23.4552 | -102.4608
23.4589 | -102.4577
23.4626 | -102.4561
23.4663 | -102.4538
23.47 | -102.4518
23.4736 | -102.4478
23.4773 | -102.4445
23.481 | -102.4427
23.4847 | -102.4386
23.4884 | -102.4363
23.4921 | -102.4363
23.4958 | -102.4368
23.4995 | -102.4364
23.5032 | -102.4363
23.5068 | -102.4346
23.5105 | -102.433
23.5142 | -102.4312
23.5179 | -102.4279
23.5216 | -102.4266
23.5253 | -102.4262
23.529 | -102.4267
23.5327 | -102.4262
23.5364 | -102.4257
23.54 | -102.4252
23.5437 | -102.4228
23.5474 | -102.422
23.5511 | -102.4205
23.5548 | -102.4181
23.5585 | -102.4164
23.5622 | -102.4153
23.5659 | -102.4125
23.5696 | -102.4109
23.5732 | -102.4075
23.5769 | -102.4035
23.5806 | -102.4005
23.5843 | -102.3965
23.588 | -102.3918
23.5917 | -102.3912
23.5954 | -102.3911
23.5991 | -102.3887
23.6027 | -102.3852
23.6064 | -102.3836
23.6101 | -102.3788
23.6138 | -102.3745
23.6175 | -102.3712
23.6212 | -102.3644
23.6249 | -102.3591
23.6286 | -102.3561
23.6323 | -102.3527
23.6359 | -102.3482
23.6396 | -102.3457
23.6433 | -102.3438
23.647 | -102.3414
23.6507 | -102.3367
23.6544 | -102.335
23.6581 | -102.332
23.6618 | -102.328
23.6655 | -102.323
23.6691 | -102.3172
23.6728 | -102.311
23.6765 | -102.3035
23.6802 | -102.2966
23.6839 | -102.289
23.6876 | -102.2795
23.6913 | -102.2691
23.695 | -102.2609
23.6986 | -102.2516
23.7023 | -102.2413
23.706 | -102.2319
23.7097 | -102.218
23.7134 | -102.2035
23.7171 | -102.191
23.7208 | -102.1771
23.7245 | -102.1599
23.7282 | -102.1438
23.7318 | -102.1278
23.7355 | -102.1123
23.7392 | -102.0986
23.7429 | -102.0843
23.7466 | -102.0691
23.7503 | -102.0547
23.754 | -102.0387
23.7577 | -102.0174
23.7614 | -101.9975
23.765 | -101.9794
23.7687 | -101.9608
23.7724 | -101.9414
23.7761 | -101.9229
23.7798 | -101.9061
23.7835 | -101.8883
23.7872 | -101.8711
23.7909 | -101.8543
23.7946 | -101.8348
23.7982 | -101.8162
23.8019 | -101.7985
23.8056 | -101.782
23.8093 | -101.7662
23.813 | -101.749
23.8167 | -101.7314
23.8204 | -101.7188
23.8241 | -101.7043
23.8277 | -101.6885
23.8314 | -101.673
23.8351 | -101.658
23.8388 | -101.6429
23.8425 | -101.6303
23.8462 | -101.6183
23.8499 | -101.6067
23.8536 | -101.5953
23.8573 | -101.5846
23.8609 | -101.5741
23.8646 | -101.5645
23.8683 | -101.5551
23.872 | -101.5463
23.8757 | -101.537
23.8794 | -101.5306
23.8831 | -101.5247
23.8868 | -101.5184
23.8905 | -101.5128
23.8941 | -101.5073
23.8978 | -101.503
23.9015 | -101.4995
23.9052 | -101.4965
23.9089 | -101.495
23.9126 | -101.4934
23.9163 | -101.4914
23.92 | -101.4887
23.9236 | -101.4863
23.9273 | -101.483
23.931 | -101.4795
23.9347 | -101.4787
23.9384 | -101.4768
23.9421 | -101.4746
23.9458 | -101.4731
23.9495 | -101.4728
23.9532 | -101.4726
23.9568 | -101.4718
23.9605 | -101.4729
23.9642 | -101.4716
23.9679 | -101.4695
23.9716 | -101.4694
23.9753 | -101.4701
23.979 | -101.4684
23.9827 | -101.4665
23.9864 | -101.4637
23.99 | -101.4582
23.9937 | -101.4533
23.9974 | -101.4505
24.0011 | -101.4485
24.0048 | -101.4475
24.0085 | -101.4472
24.0122 | -101.4461
24.0159 | -101.4441
24.0195 | -101.441
24.0232 | -101.4405
24.0269 | -101.4404
24.0306 | -101.4405
24.0343 | -101.4392
24.038 | -101.4379
24.0417 | -101.436
24.0454 | -101.4323
24.0491 | -101.4295
24.0527 | -101.4265
24.0564 | -101.4218
24.0601 | -101.4186
24.0638 | -101.4194
24.0675 | -101.4171
24.0712 | -101.4128
24.0749 | -101.408
24.0786 | -101.4024
24.0823 | -101.3976
24.0859 | -101.3947
24.0896 | -101.3893
24.0933 | -101.3852
24.097 | -101.3844
24.1007 | -101.3818
24.1044 | -101.3772
24.1081 | -101.372
24.1118 | -101.3687
24.1154 | -101.3653
24.1191 | -101.3624
24.1228 | -101.3592
24.1265 | -101.355
24.1302 | -101.3508
24.1339 | -101.3478
24.1376 | -101.3441
24.1413 | -101.3422
24.145 | -101.3392
24.1486 | -101.3377
24.1523 | -101.3379
24.156 | -101.3368
24.1597 | -101.3334
24.1634 | -101.3317
24.1671 | -101.3325
24.1708 | -101.3311
24.1745 | -101.3289
24.1782 | -101.33
24.1818 | -101.3324
24.1855 | -101.3336
24.1892 | -101.3356
24.1929 | -101.3351
24.1966 | -101.3346
24.2003 | -101.337
24.204 | -101.3378
24.2077 | -101.3379
24.2114 | -101.3406
24.215 | -101.343
24.2187 | -101.3419
24.2224 | -101.3424
24.2261 | -101.3441
24.2298 | -101.345
24.2335 | -101.3463
24.2372 | -101.3472
24.2409 | -101.3464
24.2445 | -101.3461
24.2482 | -101.3457
24.2519 | -101.3439
24.2556 | -101.3409
24.2593 | -101.3402
24.263 | -101.3386
24.2667 | -101.3347
24.2704 | -101.3261
24.2741 | -101.3055
24.2777 | -101.2742
24.2814 | -101.2321
24.2851 | -101.1923
24.2888 | -101.1656
24.2925 | -101.1415
24.2962 | -101.114
24.2999 | -101.0843
24.3036 | -101.054
24.3073 | -101.0209
24.3109 | -100.9858
24.3146 | -100.9484
24.3183 | -100.9079
24.322 | -100.8661
24.3257 | -100.824
24.3294 | -100.7812
24.3331 | -100.739
24.3368 | -100.693
24.3404 | -100.6426
24.3441 | -100.5921
24.3478 | -100.5412
24.3515 | -100.4901
24.3552 | -100.4398
24.3589 | -100.3883
24.3626 | -100.3352
24.3663 | -100.2814
24.37 | -100.229
24.3736 | -100.1773
24.3773 | -100.1267
24.381 | -100.0777
24.3847 | -100.0295
24.3884 | -99.9825
24.3921 | -99.9395
24.3958 | -99.8975
24.3995 | -99.8576
24.4032 | -99.8215
24.4068 | -99.7862
24.4105 | -99.7528
24.4142 | -99.7244
24.4179 | -99.6984
24.4216 | -99.6734
24.4253 | -99.6532
24.429 | -99.6364
24.4327 | -99.6234
24.4364 | -99.6139
24.44 | -99.6067
24.4437 | -99.6022
24.4474 | -99.6016
24.4511 | -99.6048
24.4548 | -99.6085
24.4585 | -99.6163
24.4622 | -99.6273
24.4659 | -99.6429
24.4695 | -99.6637
24.4732 | -99.6873
24.4769 | -99.7322
24.4806 | -99.7585
24.4843 | -99.7874
24.488 | -99.818
24.4917 | -99.8516
24.4954 | -99.8871
24.4991 | -99.924
24.5027 | -99.9622
24.5064 | -100.002
24.5101 | -100.0435
24.5138 | -100.0853
24.5175 | -100.1276
24.5212 | -100.1706
24.5249 | -100.2128
24.5286 | -100.253
24.5323 | -100.2952
24.5359 | -100.3381
24.5396 | -100.3822
24.5433 | -100.4249
24.547 | -100.4629
24.5507 | -100.4974
24.5544 | -100.5296
24.5581 | -100.5608
24.5618 | -100.5893
24.5654 | -100.6157
24.5691 | -100.6378
24.5728 | -100.6556
24.5765 | -100.6629
24.5802 | -100.6482
24.5839 | -100.6087
24.5876 | -100.5616
24.5913 | -100.5318
24.595 | -100.5074
24.5986 | -100.4809
24.6023 | -100.4515
24.606 | -100.4172
24.6097 | -100.3796
24.6134 | -100.3384
24.6171 | -100.296
24.6208 | -100.2558
24.6245 | -100.2177
24.6282 | -100.1786
24.6318 | -100.1268
24.6355 | -100.0582
24.6392 | -99.984
24.6429 | -99.9079
24.6466 | -99.8302
24.6503 | -99.7503
24.654 | -99.6706
24.6577 | -99.5899
24.6613 | -99.5075
24.665 | -99.4249
24.6687 | -99.3428
24.6724 | -99.2606
24.6761 | -99.1781
24.6798 | -99.0962
24.6835 | -99.0174
24.6872 | -98.9399
24.6909 | -98.8634
24.6945 | -98.7902
24.6982 | -98.7186
24.7019 | -98.6496
24.7056 | -98.5841
24.7093 | -98.5232
24.713 | -98.4665
24.7167 | -98.4122
24.7204 | -98.362
24.7241 | -98.3158
24.7277 | -98.2749
24.7314 | -98.2397
24.7351 | -98.2099
24.7388 | -98.1846
24.7425 | -98.1637
24.7462 | -98.1482
24.7499 | -98.1376
24.7536 | -98.1319
24.7572 | -98.1318
24.7609 | -98.1364
24.7646 | -98.1466
24.7683 | -98.1636
24.772 | -98.1849
24.7757 | -98.2104
24.7794 | -98.2398
24.7831 | -98.275
24.7868 | -98.3167
24.7904 | -98.3645
24.7941 | -98.4161
24.7978 | -98.4714
24.8015 | -98.5323
24.8052 | -98.5963
24.8089 | -98.6635
24.8126 | -98.7341
24.8163 | -98.8065
24.82 | -98.8817
24.8236 | -98.9604
24.8273 | -99.042
24.831 | -99.1255
24.8347 | -99.2106
24.8384 | -99.2976
24.8421 | -99.3846
24.8458 | -99.4707
24.8495 | -99.5584
24.8532 | -99.6482
24.8568 | -99.7383
24.8605 | -99.8273
24.8642 | -99.9157
24.8679 | -100.0036
24.8716 | -100.0904
24.8753 | -100.1745
24.879 | -100.2579
24.8827 | -100.3379
24.8863 | -100.415
24.89 | -100.4888
24.8937 | -100.5604
24.8974 | -100.63
24.9011 | -100.6936
24.9048 | -100.7564
24.9085 | -100.8155
24.9122 | -100.8713
24.9159 | -100.9244
24.9195 | -100.9771
24.9232 | -101.0275
24.9269 | -101.0802
24.9306 | -101.1459
24.9343 | -101.2244
24.938 | -101.2965
24.9417 | -101.3344
24.9454 | -101.356
24.9491 | -101.3716
24.9527 | -101.3835
24.9564 | -101.3929
24.9601 | -101.399
24.9638 | -101.4049
24.9675 | -101.4098
24.9712 | -101.4131
24.9749 | -101.414
24.9786 | -101.4151
24.9822 | -101.418
24.9859 | -101.4215
24.9896 | -101.4246
24.9933 | -101.428
24.997 | -101.4316
25.0007 | -101.4334
25.0044 | -101.4356
25.0081 | -101.4408
25.0118 | -101.4419
25.0154 | -101.4441
25.0191 | -101.4468
25.0228 | -101.4492
25.0265 | -101.4533
25.0302 | -101.4591
25.0339 | -101.4645
25.0376 | -101.4677
25.0413 | -101.4735
25.045 | -101.4796
25.0486 | -101.4839
25.0523 | -101.4902
25.056 | -101.4981
25.0597 | -101.5066
25.0634 | -101.5144
25.0671 | -101.5203
25.0708 | -101.5292
25.0745 | -101.5391
25.0782 | -101.5493
25.0818 | -101.5596
25.0855 | -101.5704
25.0892 | -101.5811
25.0929 | -101.5915
25.0966 | -101.6057
25.1003 | -101.6204
25.104 | -101.6337
25.1077 | -101.6498
25.1113 | -101.6678
25.115 | -101.6852
25.1187 | -101.7032
25.1224 | -101.7216
25.1261 | -101.7396
25.1298 | -101.7579
25.1335 | -101.7793
25.1372 | -101.7988
25.1409 | -101.8163
25.1445 | -101.8321
25.1482 | -101.8485
25.1519 | -101.8646
25.1556 | -101.8819
25.1593 | -101.9019
25.163 | -101.9214
25.1667 | -101.9389
25.1704 | -101.958
25.1741 | -101.9766
25.1777 | -101.9971
25.1814 | -102.0162
25.1851 | -102.0347
25.1888 | -102.053
25.1925 | -102.068
25.1962 | -102.0833
25.1999 | -102.1026
25.2036 | -102.1191
25.2072 | -102.1334
25.2109 | -102.1473
25.2146 | -102.1579
25.2183 | -102.1665
25.222 | -102.1768
25.2257 | -102.1882
25.2294 | -102.1964
25.2331 | -102.2044
25.2368 | -102.214
25.2404 | -102.2207
25.2441 | -102.2284
25.2478 | -102.2349
25.2515 | -102.2409
25.2552 | -102.2537
25.2589 | -102.264
25.2626 | -102.268
25.2663 | -102.2737
25.27 | -102.2785
25.2736 | -102.2773
25.2773 | -102.2768
25.281 | -102.2775
25.2847 | -102.2778
25.2884 | -102.2782
25.2921 | -102.2777
25.2958 | -102.2759
25.2995 | -102.2742
25.3031 | -102.2729
25.3068 | -102.2715
25.3105 | -102.2709
25.3142 | -102.2707
25.3179 | -102.268
25.3216 | -102.2651
25.3253 | -102.2638
25.329 | -102.2631
25.3327 | -102.2613
25.3363 | -102.2596
25.34 | -102.2584
25.3437 | -102.2547
25.3474 | -102.2528
25.3511 | -102.2511
25.3548 | -102.252
25.3585 | -102.2532
25.3622 | -102.2528
25.3659 | -102.2528
25.3695 | -102.2552
25.3732 | -102.2582
25.3769 | -102.2621
25.3806 | -102.2664
25.3843 | -102.2695
25.388 | -102.2746
25.3917 | -102.2806
25.3954 | -102.2854
25.399 | -102.2905
25.4027 | -102.2966
25.4064 | -102.3027
25.4101 | -102.3084
25.4138 | -102.3126
25.4175 | -102.3168
25.4212 | -102.3228
25.4249 | -102.3273
25.4286 | -102.3291
25.4322 | -102.3305
25.4359 | -102.3336
25.4396 | -102.3361
25.4433 | -102.3366
25.447 | -102.3381
25.4507 | -102.3394
25.4544 | -102.3394
25.4581 | -102.3402
25.4618 | -102.3409
25.4654 | -102.3393
25.4691 | -102.3373
25.4728 | -102.335
25.4765 | -102.3324
25.4802 | -102.3293
25.4839 | -102.3279
25.4876 | -102.3243
25.4913 | -102.3201
25.495 | -102.3147
25.4986 | -102.3089
25.5023 | -102.3048
25.506 | -102.2984
25.5097 | -102.2931
25.5134 | -102.2851
25.5171 | -102.2748
25.5208 | -102.2659
25.5245 | -102.2569
25.5281 | -102.2476
25.5318 | -102.2405
25.5355 | -102.2302
25.5392 | -102.2252
25.5429 | -102.2178
25.5466 | -102.2104
25.5503 | -102.2034
25.554 | -102.1976
25.5577 | -102.1942
25.5613 | -102.1883
25.565 | -102.1833
25.5687 | -102.1789
25.5724 | -102.1748
25.5761 | -102.1721
25.5798 | -102.1727
25.5835 | -102.1731
25.5872 | -102.1693
25.5909 | -102.1695
25.5945 | -102.171
25.5982 | -102.171
25.6019 | -102.174
25.6056 | -102.1755
25.6093 | -102.1766
25.613 | -102.1797
25.6167 | -102.1836
25.6204 | -102.1868
25.624 | -102.1911
25.6277 | -102.1954
25.6314 | -102.1998
25.6351 | -102.2039
25.6388 | -102.2071
25.6425 | -102.2126
25.6462 | -102.2201
25.6499 | -102.2266
25.6536 | -102.2326
25.6572 | -102.2376
25.6609 | -102.244
25.6646 | -102.2513
25.6683 | -102.2567
25.672 | -102.2638
25.6757 | -102.2706
25.6794 | -102.2764
25.6831 | -102.2832
25.6868 | -102.2892
25.6904 | -102.2959
25.6941 | -102.3023
25.6978 | -102.3069
25.7015 | -102.3106
25.7052 | -102.3151
25.7089 | -102.318
25.7126 | -102.3211
25.7163 | -102.3236
25.72 | -102.3244
25.7236 | -102.3264
25.7273 | -102.3287
25.731 | -102.3297
25.7347 | -102.3287
25.7384 | -102.3266
25.7421 | -102.3236
25.7458 | -102.3173
25.7495 | -102.3082
25.7531 | -102.3052
25.7568 | -102.3029
25.7605 | -102.2989
25.7642 | -102.2934
25.7679 | -102.2851
25.7716 | -102.2768
25.7753 | -102.2717
25.779 | -102.2664
25.7827 | -102.2579
25.7863 | -102.2501
25.79 | -102.2453
25.7937 | -102.2382
25.7974 | -102.2282
25.8011 | -102.2192
25.8048 | -102.2113
25.8085 | -102.2029
25.8122 | -102.1919
25.8159 | -102.1814
25.8195 | -102.1706
25.8232 | -102.1566
25.8269 | -102.1401
25.8306 | -102.1249
25.8343 | -102.1085
25.838 | -102.0919
25.8417 | -102.0774
25.8454 | -102.0628
25.849 | -102.0512
25.8527 | -102.0388
25.8564 | -102.026
25.8601 | -102.0124
25.8638 | -101.9979
25.8675 | -101.983
25.8712 | -101.9652
25.8749 | -101.9498
25.8786 | -101.935
25.8822 | -101.9183
25.8859 | -101.9018
25.8896 | -101.8858
25.8933 | -101.8674
25.897 | -101.851
25.9007 | -101.8334
25.9044 | -101.8124
25.9081 | -101.7925
25.9118 | -101.7735
25.9154 | -101.7559
25.9191 | -101.7389
25.9228 | -101.7207
25.9265 | -101.7033
25.9302 | -101.6852
25.9339 | -101.6667
25.9376 | -101.6481
25.9413 | -101.6297
25.9449 | -101.6117
25.9486 | -101.5942
25.9523 | -101.5796
25.956 | -101.5641
25.9597 | -101.5483
25.9634 | -101.5325
25.9671 | -101.5156
25.9708 | -101.5008
25.9745 | -101.4876
25.9781 | -101.4745
25.9818 | -101.4601
25.9855 | -101.4445
25.9892 | -101.4294
25.9929 | -101.4145
25.9966 | -101.4011
26.0003 | -101.3899
26.004 | -101.3804
26.0077 | -101.3687
26.0113 | -101.3576
26.015 | -101.3469
26.0187 | -101.337
26.0224 | -101.3301
26.0261 | -101.3216
26.0298 | -101.3139
26.0335 | -101.3086
26.0372 | -101.302
26.0408 | -101.295
26.0445 | -101.289
26.0482 | -101.2827
26.0519 | -101.2766
26.0556 | -101.2721
26.0593 | -101.2688
26.063 | -101.2674
26.0667 | -101.2675
26.0704 | -101.2678
26.074 | -101.2693
26.0777 | -101.2686
26.0814 | -101.2676
26.0851 | -101.2678
26.0888 | -101.2692
26.0925 | -101.2694
26.0962 | -101.2698
26.0999 | -101.2697
26.1036 | -101.2699
26.1072 | -101.2732
26.1109 | -101.2769
26.1146 | -101.2804
26.1183 | -101.2851
26.122 | -101.2865
26.1257 | -101.2893
26.1294 | -101.2907
26.1331 | -101.2916
26.1368 | -101.2951
26.1404 | -101.2995
26.1441 | -101.3007
26.1478 | -101.301
26.1515 | -101.3034
26.1552 | -101.3041
26.1589 | -101.3026
26.1626 | -101.303
26.1663 | -101.3042
26.1699 | -101.3036
26.1736 | -101.3011
26.1773 | -101.2996
26.181 | -101.3006
26.1847 | -101.2989
26.1884 | -101.2976
26.1921 | -101.2963
26.1958 | -101.2967
26.1995 | -101.2971
26.2031 | -101.296
26.2068 | -101.2922
26.2105 | -101.2867
26.2142 | -101.2826
26.2179 | -101.2775
26.2216 | -101.2722
26.2253 | -101.2655
26.229 | -101.256
26.2327 | -101.2475
26.2363 | -101.2404
26.24 | -101.235
26.2437 | -101.2266
26.2474 | -101.2186
26.2511 | -101.2119
26.2548 | -101.2041
26.2585 | -101.1955
26.2622 | -101.1868
26.2658 | -101.1768
26.2695 | -101.1661
26.2732 | -101.1574
26.2769 | -101.148
26.2806 | -101.1357
26.2843 | -101.1212
26.288 | -101.104
26.2917 | -101.0899
26.2954 | -101.0747
26.299 | -101.0598
26.3027 | -101.0467
26.3064 | -101.0336
26.3101 | -101.0209
26.3138 | -101.0066
26.3175 | -100.9918
26.3212 | -100.9757
26.3249 | -100.9595
26.3286 | -100.944
26.3322 | -100.9268
26.3359 | -100.9123
26.3396 | -100.8973
26.3433 | -100.8792
26.347 | -100.8616
26.3507 | -100.8435
26.3544 | -100.8273
26.3581 | -100.8125
26.3618 | -100.7974
26.3654 | -100.7783
26.3691 | -100.7587
26.3728 | -100.7416
26.3765 | -100.7256
26.3802 | -100.7094
26.3839 | -100.6925
26.3876 | -100.6753
26.3913 | -100.6579
26.3949 | -100.6404
26.3986 | -100.6203
26.4023 | -100.6017
26.406 | -100.5847
26.4097 | -100.5702
26.4134 | -100.5566
26.4171 | -100.5436
26.4208 | -100.5274
26.4245 | -100.5099
26.4281 | -100.4951
26.4318 | -100.4805
26.4355 | -100.4659
26.4392 | -100.4509
26.4429 | -100.4379
26.4466 | -100.426
26.4503 | -100.4134
26.454 | -100.4009
26.4577 | -100.3859
26.4613 | -100.3724
26.465 | -100.359
26.4687 | -100.3468
26.4724 | -100.336
26.4761 | -100.3257
26.4798 | -100.312
26.4835 | -100.3004
26.4872 | -100.2904
26.4908 | -100.28
26.4945 | -100.2707
26.4982 | -100.2605
26.5019 | -100.2524
26.5056 | -100.2444
26.5093 | -100.2367
26.513 | -100.2272
26.5167 | -100.219
26.5204 | -100.2131
26.524 | -100.2067
26.5277 | -100.2039
26.5314 | -100.2018
26.5351 | -100.199
26.5388 | -100.1962
26.5425 | -100.1934
26.5462 | -100.1933
26.5499 | -100.1923
26.5536 | -100.1914
26.5572 | -100.191
26.5609 | -100.1894
26.5646 | -100.1888
26.5683 | -100.187
26.572 | -100.1848
26.5757 | -100.1853
26.5794 | -100.1854
26.5831 | -100.1845
26.5867 | -100.1838
26.5904 | -100.1825
26.5941 | -100.1828
26.5978 | -100.1849
26.6015 | -100.1881
26.6052 | -100.1899
26.6089 | -100.1907
26.6126 | -100.1873
26.6163 | -100.1866
26.6199 | -100.1873
26.6236 | -100.1871
26.6273 | -100.1857
26.631 | -100.1848
26.6347 | -100.1851
26.6384 | -100.1864
26.6421 | -100.1855
26.6458 | -100.1842
26.6495 | -100.1833
26.6531 | -100.181
26.6568 | -100.1801
26.6605 | -100.1798
26.6642 | -100.1822
26.6679 | -100.1812
26.6716 | -100.18
26.6753 | -100.1791
26.679 | -100.1739
26.6826 | -100.1701
26.6863 | -100.1656
26.69 | -100.1572
26.6937 | -100.1506
26.6974 | -100.1456
26.7011 | -100.1379
26.7048 | -100.1323
26.7085 | -100.1285
26.7122 | -100.1218
26.7158 | -100.114
26.7195 | -100.1062
26.7232 | -100.0993
26.7269 | -100.0932
26.7306 | -100.0854
26.7343 | -100.0738
26.738 | -100.0615
26.7417 | -100.0542
26.7454 | -100.0449
26.749 | -100.0322
26.7527 | -100.0215
26.7564 | -100.0103
26.7601 | -100.0002
26.7638 | -99.9905
26.7675 | -99.9796
26.7712 | -99.9678
26.7749 | -99.9592
26.7786 | -99.9446
26.7822 | -99.9318
26.7859 | -99.9198
26.7896 | -99.9062
26.7933 | -99.8922
26.797 | -99.8826
26.8007 | -99.873
26.8044 | -99.8611
26.8081 | -99.8488
26.8117 | -99.8355
26.8154 | -99.8227
26.8191 | -99.8117
26.8228 | -99.7991
26.8265 | -99.7861
26.8302 | -99.7741
26.8339 | -99.7649
26.8376 | -99.7554
26.8413 | -99.7482
26.8449 | -99.7403
26.8486 | -99.7326
26.8523 | -99.7265
26.856 | -99.7216
26.8597 | -99.7125
26.8634 | -99.706
26.8671 | -99.7014
26.8708 | -99.6953
26.8745 | -99.6902
26.8781 | -99.6858
26.8818 | -99.6832
26.8855 | -99.6811
26.8892 | -99.6805
26.8929 | -99.6788
26.8966 | -99.6774
26.9003 | -99.6775
26.904 | -99.675
26.9076 | -99.6747
26.9113 | -99.6775
26.915 | -99.6798
26.9187 | -99.6838
26.9224 | -99.6881
26.9261 | -99.6919
26.9298 | -99.6973
26.9335 | -99.7054
26.9372 | -99.7107
26.9408 | -99.7169
26.9445 | -99.7259
26.9482 | -99.7349
26.9519 | -99.7428
26.9556 | -99.7523
26.9593 | -99.7614
26.963 | -99.7681
26.9667 | -99.778
26.9704 | -99.7883
26.974 | -99.7981
26.9777 | -99.8087
26.9814 | -99.8175
26.9851 | -99.8238
26.9888 | -99.833
26.9925 | -99.8463
26.9962 | -99.8563
26.9999 | -99.8648
27.0036 | -99.871
27.0072 | -99.8757
27.0109 | -99.8822
27.0146 | -99.8862
27.0183 | -99.8861
27.022 | -99.8913
27.0257 | -99.8981
27.0294 | -99.9022
27.0331 | -99.9047
27.0367 | -99.9056
27.0404 | -99.9051
27.0441 | -99.9052
27.0478 | -99.9042
27.0515 | -99.9024
27.0552 | -99.9008
27.0589 | -99.901
27.0626 | -99.898
27.0663 | -99.8947
27.0699 | -99.8926
27.0736 | -99.884
27.0773 | -99.877
27.081 | -99.8725
27.0847 | -99.8621
27.0884 | -99.8526
27.0921 | -99.8447
27.0958 | -99.8381
27.0995 | -99.8329
27.1031 | -99.8255
27.1068 | -99.8155
27.1105 | -99.8066
27.1142 | -99.802
27.1179 | -99.7955
27.1216 | -99.7843
27.1253 | -99.776
27.129 | -99.7684
27.1326 | -99.7595
27.1363 | -99.7564
27.14 | -99.7492
27.1437 | -99.7421
27.1474 | -99.7355
27.1511 | -99.7286
27.1548 | -99.7249
27.1585 | -99.7209
27.1622 | -99.7218
27.1658 | -99.7198
27.1695 | -99.7157
27.1732 | -99.7125
27.1769 | -99.7117
27.1806 | -99.7104
27.1843 | -99.7092
27.188 | -99.7104
27.1917 | -99.7141
27.1954 | -99.7185
27.199 | -99.7232
27.2027 | -99.7259
27.2064 | -99.7309
27.2101 | -99.7381
27.2138 | -99.7468
27.2175 | -99.7544
27.2212 | -99.7615
27.2249 | -99.7693
27.2285 | -99.7777
27.2322 | -99.7873
27.2359 | -99.7977
27.2396 | -99.8102
27.2433 | -99.8258
27.247 | -99.839
27.2507 | -99.8505
27.2544 | -99.8637
27.2581 | -99.879
27.2617 | -99.8937
27.2654 | -99.9089
27.2691 | -99.9256
27.2728 | -99.9419
27.2765 | -99.9594
27.2802 | -99.9777
27.2839 | -99.995
27.2876 | -100.0147
27.2913 | -100.0339
27.2949 | -100.0513
27.2986 | -100.071
27.3023 | -100.0908
27.306 | -100.1096
27.3097 | -100.1271
27.3134 | -100.1438
27.3171 | -100.1619
27.3208 | -100.1838
27.3244 | -100.2061
27.3281 | -100.2264
27.3318 | -100.2448
27.3355 | -100.2644
27.3392 | -100.2843
27.3429 | -100.3038
27.3466 | -100.3238
27.3503 | -100.3412
27.354 | -100.3583
27.3576 | -100.3743
27.3613 | -100.3912
27.365 | -100.4079
27.3687 | -100.4256
27.3724 | -100.442
27.3761 | -100.4566
27.3798 | -100.4718
27.3835 | -100.4839
27.3872 | -100.4947
27.3908 | -100.5087
27.3945 | -100.5215
27.3982 | -100.5346
27.4019 | -100.5503
27.4056 | -100.5623
27.4093 | -100.5699
27.413 | -100.5767
27.4167 | -100.5855
27.4204 | -100.5926
27.424 | -100.5993
27.4277 | -100.604
27.4314 | -100.6081
27.4351 | -100.6111
27.4388 | -100.6135
27.4425 | -100.6156
27.4462 | -100.6173
27.4499 | -100.6176
27.4535 | -100.621
27.4572 | -100.6277
27.4609 | -100.6284
27.4646 | -100.6272
27.4683 | -100.6277
27.472 | -100.6253
27.4757 | -100.6213
27.4794 | -100.6177
27.4831 | -100.6138
27.4867 | -100.6082
27.4904 | -100.6039
27.4941 | -100.5992
27.4978 | -100.5908
27.5015 | -100.5846
27.5052 | -100.581
27.5089 | -100.5745
27.5126 | -100.5687
27.5163 | -100.5628
27.5199 | -100.5565
27.5236 | -100.5506
27.5273 | -100.5448
27.531 | -100.54
27.5347 | -100.5355
27.5384 | -100.5331
27.5421 | -100.5287
27.5458 | -100.5246
27.5494 | -100.5205
27.5531 | -100.5167
27.5568 | -100.5148
27.5605 | -100.5153
27.5642 | -100.5142
27.5679 | -100.5112
27.5716 | -100.5103
27.5753 | -100.5094
27.579 | -100.508
27.5826 | -100.5087
27.5863 | -100.5117
27.59 | -100.5155
27.5937 | -100.5187
27.5974 | -100.5219
27.6011 | -100.522
27.6048 | -100.5237
27.6085 | -100.5251
27.6122 | -100.5251
27.6158 | -100.5263
27.6195 | -100.5272
27.6232 | -100.5289
27.6269 | -100.5298
27.6306 | -100.5324
27.6343 | -100.5365
27.638 | -100.5419
27.6417 | -100.546
27.6454 | -100.5488
27.649 | -100.5513
27.6527 | -100.552
27.6564 | -100.551
27.6601 | -100.5496
27.6638 | -100.5502
27.6675 | -100.5416
27.6712 | -100.5388
27.6749 | -100.5353
27.6785 | -100.53
27.6822 | -100.5235
27.6859 | -100.5179
27.6896 | -100.5096
27.6933 | -100.4997
27.697 | -100.4894
27.7007 | -100.4815
27.7044 | -100.4729
27.7081 | -100.4604
27.7117 | -100.4459
27.7154 | -100.4308
27.7191 | -100.4152
27.7228 | -100.3987
27.7265 | -100.3802
27.7302 | -100.3625
27.7339 | -100.3457
27.7376 | -100.3283
27.7413 | -100.309
27.7449 | -100.2895
27.7486 | -100.2707
27.7523 | -100.2507
27.756 | -100.2319
27.7597 | -100.213
27.7634 | -100.1935
27.7671 | -100.1728
27.7708 | -100.1516
27.7744 | -100.1335
27.7781 | -100.1177
27.7818 | -100.0999
27.7855 | -100.0806
27.7892 | -100.061
27.7929 | -100.0411
27.7966 | -100.0207
27.8003 | -100.0032
27.804 | -99.989
27.8076 | -99.9757
27.8113 | -99.962
27.815 | -99.9477
27.8187 | -99.9344
27.8224 | -99.9214
27.8261 | -99.9098
27.8298 | -99.8985
27.8335 | -99.8885
27.8372 | -99.8791
27.8408 | -99.8729
27.8445 | -99.8662
27.8482 | -99.8617
27.8519 | -99.8587
27.8556 | -99.8534
27.8593 | -99.8472
27.863 | -99.845
27.8667 | -99.8461
27.8703 | -99.8472
27.874 | -99.8487
27.8777 | -99.8532
27.8814 | -99.8593
27.8851 | -99.8662
27.8888 | -99.8743
27.8925 | -99.8833
27.8962 | -99.8918
27.8999 | -99.9011
27.9035 | -99.9098
27.9072 | -99.9184
27.9109 | -99.9275
27.9146 | -99.9356
27.9183 | -99.9449
27.922 | -99.9542
27.9257 | -99.9578
27.9294 | -99.9546
27.9331 | -99.9571
27.9367 | -99.9635
27.9404 | -99.9686
27.9441 | -99.9753
27.9478 | -99.9826
27.9515 | -99.9883
27.9552 | -99.9945
27.9589 | -100.0027
27.9626 | -100.0097
27.9662 | -100.015
27.9699 | -100.0184
27.9736 | -100.0196
27.9773 | -100.0203
27.981 | -100.017
27.9847 | -100.0115
27.9884 | -100.008
27.9921 | -100.006
27.9958 | -100.0029
27.9994 | -99.9976
28.0031 | -99.9928