## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 263 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 273
Hood | 02-92B/P | 213
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1013.1 | hPa |  | PRC PRO 01
Humidity | 56.2 | %/Hr
Internal temperature of calibration bench | 24.37 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1021 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.8 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.97 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.99 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3062 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 10/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 11/07/2018
Technical Manager | Anthony HUE | Approval |  | 16/07/2017
CEO | Robert GLEMOT | Approval |  | 16/07/2017
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3999 | 57.367
0.012209 | -35.9516 | 38.2941
0.018313 | -34.9042 | 28.3003
0.024417 | -34.455 | 21.9472
0.030521 | -34.3504 | 18.1748
0.036626 | -34.2104 | 15.0874
0.04273 | -34.1264 | 12.7688
0.048834 | -34.0771 | 11.2031
0.054938 | -34.0089 | 10.0104
0.061043 | -33.9382 | 8.8917
0.067147 | -33.9286 | 8.0286
0.073251 | -33.9301 | 7.2723
0.079355 | -33.9486 | 6.6955
0.08546 | -33.9667 | 6.2152
0.091564 | -33.964 | 5.8314
0.097668 | -33.9337 | 5.5786
0.10377 | -33.8911 | 5.2507
0.10988 | -33.8632 | 4.9045
0.11598 | -33.8544 | 4.5572
0.12209 | -33.8689 | 4.277
0.12819 | -33.898 | 4.0344
0.13429 | -33.9264 | 3.7859
0.1404 | -33.9415 | 3.5145
0.1465 | -33.9551 | 3.2588
0.15261 | -33.9611 | 3.0249
0.15871 | -33.9578 | 2.8363
0.16482 | -33.9454 | 2.688
0.17092 | -33.9254 | 2.6292
0.17702 | -33.91 | 2.5357
0.18313 | -33.896 | 2.4532
0.18923 | -33.8859 | 2.3463
0.19534 | -33.8802 | 2.2387
0.20144 | -33.8794 | 2.1207
0.20754 | -33.8814 | 2.039
0.21365 | -33.8835 | 1.9576
0.21975 | -33.8872 | 1.8874
0.22586 | -33.8942 | 1.8187
0.23196 | -33.9019 | 1.7465
0.23807 | -33.9098 | 1.6721
0.24417 | -33.9124 | 1.5859
0.25027 | -33.916 | 1.5064
0.25638 | -33.9166 | 1.4205
0.26248 | -33.9155 | 1.3379
0.26859 | -33.914 | 1.2584
0.27469 | -33.9137 | 1.1751
0.2808 | -33.9189 | 1.1093
0.2869 | -33.9229 | 1.0407
0.293 | -33.9278 | 0.98198
0.29911 | -33.9333 | 0.93989
0.30521 | -33.9386 | 0.89948
0.31132 | -33.9422 | 0.88139
0.31742 | -33.9462 | 0.85566
0.32353 | -33.9492 | 0.83456
0.32963 | -33.952 | 0.81699
0.33573 | -33.9539 | 0.80253
0.34184 | -33.9564 | 0.78511
0.34794 | -33.9596 | 0.75168
0.35405 | -33.9628 | 0.72254
0.36015 | -33.9662 | 0.6903
0.36626 | -33.9686 | 0.6548
0.37236 | -33.9711 | 0.60723
0.37846 | -33.9733 | 0.55236
0.38457 | -33.9727 | 0.49406
0.39067 | -33.9735 | 0.44468
0.39678 | -33.9723 | 0.38947
0.40288 | -33.9705 | 0.33649
0.40899 | -33.9675 | 0.28508
0.41509 | -33.9637 | 0.23354
0.42119 | -33.9598 | 0.18738
0.4273 | -33.9568 | 0.13933
0.4334 | -33.9541 | 0.090622
0.43951 | -33.9517 | 0.043063
0.44561 | -33.9508 | -0.0012692
0.45172 | -33.9524 | -0.037663
0.45782 | -33.9525 | -0.07848
0.46392 | -33.9538 | -0.11631
0.47003 | -33.9555 | -0.15064
0.47613 | -33.9575 | -0.18271
0.48224 | -33.9595 | -0.21361
0.48834 | -33.9615 | -0.23704
0.49445 | -33.9633 | -0.26328
0.50055 | -33.9654 | -0.2871
0.50665 | -33.9669 | -0.30779
0.51276 | -33.9678 | -0.32972
0.51886 | -33.9687 | -0.34954
0.52497 | -33.9679 | -0.35878
0.53107 | -33.9682 | -0.37202
0.53717 | -33.9685 | -0.38583
0.54328 | -33.9683 | -0.39376
0.54938 | -33.9684 | -0.40406
0.55549 | -33.9681 | -0.41483
0.56159 | -33.9683 | -0.4258
0.5677 | -33.9681 | -0.43679
0.5738 | -33.968 | -0.45032
0.5799 | -33.9682 | -0.46274
0.58601 | -33.9688 | -0.47798
0.59211 | -33.9701 | -0.49773
0.59822 | -33.9709 | -0.51316
0.60432 | -33.9721 | -0.53304
0.61043 | -33.9732 | -0.55199
0.61653 | -33.9745 | -0.57156
0.62263 | -33.9757 | -0.59229
0.62874 | -33.9771 | -0.61094
0.63484 | -33.9782 | -0.63001
0.64095 | -33.9794 | -0.64921
0.64705 | -33.9805 | -0.66745
0.65316 | -33.9815 | -0.68579
0.65926 | -33.9825 | -0.70712
0.66536 | -33.9819 | -0.72524
0.67147 | -33.9822 | -0.74101
0.67757 | -33.9826 | -0.7583
0.68368 | -33.9825 | -0.77802
0.68978 | -33.9825 | -0.79581
0.69589 | -33.9821 | -0.81413
0.70199 | -33.9814 | -0.83595
0.70809 | -33.9809 | -0.85582
0.7142 | -33.9805 | -0.87514
0.7203 | -33.9797 | -0.89448
0.72641 | -33.979 | -0.91541
0.73251 | -33.9788 | -0.93586
0.73862 | -33.9789 | -0.9595
0.74472 | -33.9791 | -0.97877
0.75082 | -33.9788 | -0.99982
0.75693 | -33.9789 | -1.0228
0.76303 | -33.9791 | -1.045
0.76914 | -33.9799 | -1.0708
0.77524 | -33.9803 | -1.0947
0.78135 | -33.9809 | -1.1173
0.78745 | -33.9817 | -1.1412
0.79355 | -33.9823 | -1.1649
0.79966 | -33.9835 | -1.1895
0.80576 | -33.9849 | -1.218
0.81187 | -33.9855 | -1.246
0.81797 | -33.9864 | -1.275
0.82408 | -33.9872 | -1.301
0.83018 | -33.988 | -1.3303
0.83628 | -33.9891 | -1.3587
0.84239 | -33.9893 | -1.3854
0.84849 | -33.9896 | -1.4123
0.8546 | -33.9902 | -1.4385
0.8607 | -33.9905 | -1.4661
0.86681 | -33.9909 | -1.4928
0.87291 | -33.9907 | -1.5172
0.87901 | -33.9905 | -1.5386
0.88512 | -33.9905 | -1.5607
0.89122 | -33.9906 | -1.583
0.89733 | -33.9905 | -1.6015
0.90343 | -33.9904 | -1.622
0.90953 | -33.9906 | -1.6366
0.91564 | -33.9902 | -1.6524
0.92174 | -33.99 | -1.6666
0.92785 | -33.9897 | -1.6801
0.93395 | -33.9898 | -1.692
0.94006 | -33.9899 | -1.7053
0.94616 | -33.9897 | -1.711
0.95226 | -33.9899 | -1.7225
0.95837 | -33.9901 | -1.7344
0.96447 | -33.9905 | -1.7449
0.97058 | -33.9905 | -1.7552
0.97668 | -33.9904 | -1.7639
0.98279 | -33.9904 | -1.769
0.98889 | -33.9906 | -1.7746
0.99499 | -33.9907 | -1.7838
1.0011 | -33.991 | -1.7895
1.0072 | -33.9909 | -1.7945
1.0133 | -33.9908 | -1.7994
1.0194 | -33.9908 | -1.806
1.0255 | -33.9905 | -1.8138
1.0316 | -33.99 | -1.8191
1.0377 | -33.9896 | -1.8262
1.0438 | -33.9889 | -1.8319
1.0499 | -33.9875 | -1.8397
1.056 | -33.9865 | -1.846
1.0621 | -33.9855 | -1.8531
1.0682 | -33.9851 | -1.8592
1.0743 | -33.9841 | -1.8682
1.0805 | -33.9835 | -1.8762
1.0866 | -33.982 | -1.8876
1.0927 | -33.9808 | -1.8961
1.0988 | -33.9801 | -1.9073
1.1049 | -33.9788 | -1.9159
1.111 | -33.9774 | -1.9276
1.1171 | -33.9761 | -1.9377
1.1232 | -33.975 | -1.952
1.1293 | -33.974 | -1.9642
1.1354 | -33.9732 | -1.9769
1.1415 | -33.9721 | -1.9911
1.1476 | -33.9709 | -2.0052
1.1537 | -33.9698 | -2.0186
1.1598 | -33.9697 | -2.0362
1.1659 | -33.9693 | -2.0484
1.172 | -33.9686 | -2.0617
1.1781 | -33.968 | -2.0756
1.1842 | -33.9676 | -2.0909
1.1903 | -33.9683 | -2.1053
1.1964 | -33.9682 | -2.1195
1.2025 | -33.9682 | -2.1344
1.2086 | -33.9687 | -2.1495
1.2147 | -33.9689 | -2.164
1.2209 | -33.9696 | -2.1769
1.227 | -33.9713 | -2.1912
1.2331 | -33.9722 | -2.205
1.2392 | -33.9731 | -2.22
1.2453 | -33.9742 | -2.2318
1.2514 | -33.9754 | -2.2434
1.2575 | -33.9768 | -2.2568
1.2636 | -33.9789 | -2.2677
1.2697 | -33.9805 | -2.2788
1.2758 | -33.982 | -2.2923
1.2819 | -33.9836 | -2.3051
1.288 | -33.9852 | -2.3182
1.2941 | -33.987 | -2.3284
1.3002 | -33.9888 | -2.3364
1.3063 | -33.9904 | -2.3461
1.3124 | -33.992 | -2.3558
1.3185 | -33.9938 | -2.366
1.3246 | -33.9955 | -2.3764
1.3307 | -33.9974 | -2.3855
1.3368 | -33.9991 | -2.393
1.3429 | -34.0006 | -2.4017
1.349 | -34.002 | -2.4101
1.3551 | -34.0035 | -2.4206
1.3613 | -34.005 | -2.4283
1.3674 | -34.0062 | -2.4368
1.3735 | -34.0073 | -2.445
1.3796 | -34.0085 | -2.4525
1.3857 | -34.0094 | -2.4596
1.3918 | -34.0104 | -2.4693
1.3979 | -34.0117 | -2.4777
1.404 | -34.0123 | -2.4877
1.4101 | -34.0132 | -2.4967
1.4162 | -34.014 | -2.5054
1.4223 | -34.0148 | -2.5157
1.4284 | -34.0156 | -2.5242
1.4345 | -34.0162 | -2.5344
1.4406 | -34.0166 | -2.5458
1.4467 | -34.0173 | -2.5554
1.4528 | -34.018 | -2.5648
1.4589 | -34.0188 | -2.5754
1.465 | -34.019 | -2.5835
1.4711 | -34.0199 | -2.5928
1.4772 | -34.0205 | -2.605
1.4833 | -34.0211 | -2.6172
1.4894 | -34.0218 | -2.6286
1.4955 | -34.0226 | -2.6401
1.5016 | -34.0232 | -2.652
1.5078 | -34.0237 | -2.6664
1.5139 | -34.0244 | -2.6789
1.52 | -34.0249 | -2.6919
1.5261 | -34.0253 | -2.7036
1.5322 | -34.0256 | -2.7163
1.5383 | -34.0259 | -2.7284
1.5444 | -34.0266 | -2.7415
1.5505 | -34.027 | -2.7547
1.5566 | -34.0276 | -2.7674
1.5627 | -34.0282 | -2.7807
1.5688 | -34.0289 | -2.7938
1.5749 | -34.0292 | -2.8061
1.581 | -34.0296 | -2.8221
1.5871 | -34.0302 | -2.8373
1.5932 | -34.0306 | -2.85
1.5993 | -34.031 | -2.8663
1.6054 | -34.0314 | -2.8808
1.6115 | -34.0318 | -2.8972
1.6176 | -34.0321 | -2.9128
1.6237 | -34.0324 | -2.927
1.6298 | -34.0327 | -2.9418
1.6359 | -34.0331 | -2.956
1.642 | -34.0336 | -2.9699
1.6482 | -34.034 | -2.9861
1.6543 | -34.0345 | -3.0023
1.6604 | -34.0349 | -3.0164
1.6665 | -34.035 | -3.0308
1.6726 | -34.0353 | -3.0463
1.6787 | -34.0358 | -3.0625
1.6848 | -34.0357 | -3.0782
1.6909 | -34.0361 | -3.0938
1.697 | -34.0363 | -3.1092
1.7031 | -34.0364 | -3.1243
1.7092 | -34.0364 | -3.1386
1.7153 | -34.0366 | -3.154
1.7214 | -34.0363 | -3.1688
1.7275 | -34.0362 | -3.184
1.7336 | -34.036 | -3.1996
1.7397 | -34.0361 | -3.2143
1.7458 | -34.0362 | -3.2301
1.7519 | -34.036 | -3.2455
1.758 | -34.036 | -3.2598
1.7641 | -34.0361 | -3.2737
1.7702 | -34.0359 | -3.2889
1.7763 | -34.036 | -3.3022
1.7824 | -34.0358 | -3.3177
1.7885 | -34.0357 | -3.3315
1.7947 | -34.0358 | -3.3458
1.8008 | -34.0357 | -3.3583
1.8069 | -34.0357 | -3.3716
1.813 | -34.0357 | -3.3854
1.8191 | -34.0355 | -3.3993
1.8252 | -34.0354 | -3.413
1.8313 | -34.0354 | -3.427
1.8374 | -34.0355 | -3.4402
1.8435 | -34.0356 | -3.4523
1.8496 | -34.0353 | -3.4665
1.8557 | -34.0353 | -3.4808
1.8618 | -34.0354 | -3.4927
1.8679 | -34.0354 | -3.5055
1.874 | -34.0355 | -3.5184
1.8801 | -34.0354 | -3.5312
1.8862 | -34.0356 | -3.5439
1.8923 | -34.0356 | -3.556
1.8984 | -34.0353 | -3.5669
1.9045 | -34.0353 | -3.578
1.9106 | -34.0351 | -3.5898
1.9167 | -34.0349 | -3.6012
1.9228 | -34.0349 | -3.6122
1.9289 | -34.0347 | -3.6216
1.9351 | -34.0348 | -3.6309
1.9412 | -34.0349 | -3.6414
1.9473 | -34.035 | -3.6533
1.9534 | -34.035 | -3.6645
1.9595 | -34.0349 | -3.676
1.9656 | -34.0346 | -3.6869
1.9717 | -34.0344 | -3.6976
1.9778 | -34.0345 | -3.709
1.9839 | -34.0346 | -3.7204
1.99 | -34.0347 | -3.7304
1.9961 | -34.0345 | -3.7411
2.0022 | -34.0345 | -3.7513
2.0083 | -34.0343 | -3.7614
2.0144 | -34.0343 | -3.7715
2.0205 | -34.0344 | -3.7832
2.0266 | -34.0346 | -3.7947
2.0327 | -34.0347 | -3.8059
2.0388 | -34.035 | -3.8193
2.0449 | -34.0351 | -3.8297
2.051 | -34.0352 | -3.8398
2.0571 | -34.0352 | -3.8512
2.0632 | -34.0351 | -3.8618
2.0693 | -34.0352 | -3.8731
2.0754 | -34.0354 | -3.8847
2.0816 | -34.0356 | -3.8966
2.0877 | -34.0356 | -3.9077
2.0938 | -34.0359 | -3.92
2.0999 | -34.0359 | -3.9315
2.106 | -34.0361 | -3.9433
2.1121 | -34.0362 | -3.9555
2.1182 | -34.0364 | -3.9656
2.1243 | -34.0366 | -3.9761
2.1304 | -34.0367 | -3.9864
2.1365 | -34.0371 | -3.9979
2.1426 | -34.0372 | -4.0109
2.1487 | -34.0377 | -4.0225
2.1548 | -34.0381 | -4.0352
2.1609 | -34.0387 | -4.0471
2.167 | -34.0391 | -4.0593
2.1731 | -34.0395 | -4.0712
2.1792 | -34.0401 | -4.0823
2.1853 | -34.0406 | -4.0949
2.1914 | -34.0411 | -4.1069
2.1975 | -34.0416 | -4.1188
2.2036 | -34.042 | -4.1318
2.2097 | -34.0425 | -4.1436
2.2158 | -34.0433 | -4.1538
2.222 | -34.0438 | -4.1639
2.2281 | -34.0442 | -4.1747
2.2342 | -34.0447 | -4.1875
2.2403 | -34.0451 | -4.1976
2.2464 | -34.0458 | -4.2085
2.2525 | -34.0462 | -4.2193
2.2586 | -34.047 | -4.2292
2.2647 | -34.0478 | -4.2387
2.2708 | -34.0483 | -4.2502
2.2769 | -34.0491 | -4.2615
2.283 | -34.0496 | -4.2711
2.2891 | -34.0501 | -4.2815
2.2952 | -34.0506 | -4.2914
2.3013 | -34.0512 | -4.3016
2.3074 | -34.0519 | -4.3113
2.3135 | -34.0523 | -4.3226
2.3196 | -34.0529 | -4.3314
2.3257 | -34.0532 | -4.3419
2.3318 | -34.0535 | -4.3512
2.3379 | -34.054 | -4.3624
2.344 | -34.0543 | -4.3733
2.3501 | -34.0548 | -4.384
2.3562 | -34.0551 | -4.3927
2.3623 | -34.0554 | -4.4016
2.3685 | -34.0556 | -4.4108
2.3746 | -34.0559 | -4.4214
2.3807 | -34.0563 | -4.4306
2.3868 | -34.0565 | -4.4398
2.3929 | -34.0569 | -4.4493
2.399 | -34.0572 | -4.4577
2.4051 | -34.0576 | -4.4673
2.4112 | -34.0579 | -4.4765
2.4173 | -34.0583 | -4.4857
2.4234 | -34.0583 | -4.4946
2.4295 | -34.0586 | -4.5035
2.4356 | -34.0587 | -4.513
2.4417 | -34.0588 | -4.5232
2.4478 | -34.0589 | -4.533
2.4539 | -34.0592 | -4.5412
2.46 | -34.059 | -4.5512
2.4661 | -34.0591 | -4.5607
2.4722 | -34.0593 | -4.5697
2.4783 | -34.0595 | -4.5784
2.4844 | -34.0595 | -4.5873
2.4905 | -34.0595 | -4.5968
2.4966 | -34.0593 | -4.606
2.5027 | -34.0593 | -4.6144
2.5089 | -34.0593 | -4.6233
2.515 | -34.0591 | -4.6328
2.5211 | -34.0591 | -4.6415
2.5272 | -34.0589 | -4.6508
2.5333 | -34.0588 | -4.6606
2.5394 | -34.0588 | -4.6699
2.5455 | -34.0587 | -4.6796
2.5516 | -34.0587 | -4.6886
2.5577 | -34.0587 | -4.6988
2.5638 | -34.0586 | -4.7094
2.5699 | -34.0584 | -4.7198
2.576 | -34.0583 | -4.7304
2.5821 | -34.0581 | -4.7405
2.5882 | -34.058 | -4.7504
2.5943 | -34.0579 | -4.7603
2.6004 | -34.0579 | -4.7705
2.6065 | -34.0577 | -4.7792
2.6126 | -34.0577 | -4.7886
2.6187 | -34.0577 | -4.7976
2.6248 | -34.0576 | -4.807
2.6309 | -34.0575 | -4.8183
2.637 | -34.0576 | -4.8295
2.6431 | -34.0576 | -4.8392
2.6492 | -34.0575 | -4.8494
2.6554 | -34.0576 | -4.8598
2.6615 | -34.0575 | -4.8699
2.6676 | -34.0575 | -4.8805
2.6737 | -34.0577 | -4.8913
2.6798 | -34.0577 | -4.9015
2.6859 | -34.0578 | -4.9121
2.692 | -34.0578 | -4.9221
2.6981 | -34.0578 | -4.9332
2.7042 | -34.0579 | -4.9451
2.7103 | -34.0579 | -4.9553
2.7164 | -34.0582 | -4.966
2.7225 | -34.0583 | -4.9774
2.7286 | -34.0584 | -4.9885
2.7347 | -34.0585 | -4.9986
2.7408 | -34.0588 | -5.008
2.7469 | -34.0588 | -5.0194
2.753 | -34.0592 | -5.03
2.7591 | -34.0593 | -5.0419
2.7652 | -34.0594 | -5.0534
2.7713 | -34.0594 | -5.0634
2.7774 | -34.0599 | -5.0748
2.7835 | -34.06 | -5.0855
2.7896 | -34.0602 | -5.0951
2.7958 | -34.0606 | -5.1055
2.8019 | -34.0611 | -5.1154
2.808 | -34.0614 | -5.1271
2.8141 | -34.0617 | -5.1397
2.8202 | -34.0621 | -5.1517
2.8263 | -34.0626 | -5.1626
2.8324 | -34.0628 | -5.1749
2.8385 | -34.0633 | -5.1867
2.8446 | -34.0638 | -5.1983
2.8507 | -34.0642 | -5.2101
2.8568 | -34.0647 | -5.2212
2.8629 | -34.0651 | -5.2324
2.869 | -34.0655 | -5.2436
2.8751 | -34.0659 | -5.2554
2.8812 | -34.0665 | -5.2683
2.8873 | -34.0669 | -5.2796
2.8934 | -34.0673 | -5.2914
2.8995 | -34.0676 | -5.3032
2.9056 | -34.068 | -5.314
2.9117 | -34.0684 | -5.3255
2.9178 | -34.0691 | -5.3378
2.9239 | -34.0695 | -5.3508
2.93 | -34.07 | -5.364
2.9361 | -34.0705 | -5.3753
2.9423 | -34.071 | -5.387
2.9484 | -34.0715 | -5.4001
2.9545 | -34.0722 | -5.4129
2.9606 | -34.0727 | -5.4259
2.9667 | -34.0732 | -5.4376
2.9728 | -34.0736 | -5.4497
2.9789 | -34.0741 | -5.4621
2.985 | -34.0745 | -5.4754
2.9911 | -34.075 | -5.4877
2.9972 | -34.0755 | -5.5001
3.0033 | -34.0761 | -5.5129
3.0094 | -34.0766 | -5.5256
3.0155 | -34.0772 | -5.5379
3.0216 | -34.0777 | -5.5509
3.0277 | -34.0781 | -5.5642
3.0338 | -34.0787 | -5.5774
3.0399 | -34.0793 | -5.5907
3.046 | -34.0796 | -5.6034
3.0521 | -34.0802 | -5.6156
3.0582 | -34.0805 | -5.6291
3.0643 | -34.081 | -5.6418
3.0704 | -34.0813 | -5.6549
3.0765 | -34.0816 | -5.6669
3.0827 | -34.082 | -5.6797
3.0888 | -34.0823 | -5.6922
3.0949 | -34.0827 | -5.7049
3.101 | -34.0831 | -5.7185
3.1071 | -34.0834 | -5.7321
3.1132 | -34.0839 | -5.7458
3.1193 | -34.0843 | -5.7585
3.1254 | -34.0847 | -5.7705
3.1315 | -34.0851 | -5.7843
3.1376 | -34.0854 | -5.7969
3.1437 | -34.0858 | -5.8101
3.1498 | -34.0861 | -5.8239
3.1559 | -34.0865 | -5.8368
3.162 | -34.0865 | -5.8504
3.1681 | -34.0867 | -5.8642
3.1742 | -34.0871 | -5.8775
3.1803 | -34.0874 | -5.8904
3.1864 | -34.0878 | -5.9025
3.1925 | -34.0881 | -5.9157
3.1986 | -34.0882 | -5.9283
3.2047 | -34.0884 | -5.9419
3.2108 | -34.0887 | -5.9535
3.2169 | -34.0891 | -5.9668
3.223 | -34.0891 | -5.9799
3.2292 | -34.0891 | -5.9934
3.2353 | -34.0891 | -6.0031
3.2414 | -34.0893 | -6.0158
3.2475 | -34.0893 | -6.0268
3.2536 | -34.0896 | -6.0401
3.2597 | -34.0898 | -6.0512
3.2658 | -34.09 | -6.062
3.2719 | -34.0902 | -6.073
3.278 | -34.0903 | -6.085
3.2841 | -34.0905 | -6.0968
3.2902 | -34.0906 | -6.1091
3.2963 | -34.0908 | -6.1207
3.3024 | -34.0907 | -6.131
3.3085 | -34.091 | -6.1433
3.3146 | -34.0909 | -6.1542
3.3207 | -34.0911 | -6.1662
3.3268 | -34.0913 | -6.1783
3.3329 | -34.0914 | -6.1882
3.339 | -34.0915 | -6.198
3.3451 | -34.0916 | -6.2093
3.3512 | -34.0917 | -6.219
3.3573 | -34.0919 | -6.2302
3.3634 | -34.092 | -6.2416
3.3696 | -34.0921 | -6.2534
3.3757 | -34.0921 | -6.2634
3.3818 | -34.0924 | -6.2749
3.3879 | -34.0924 | -6.2854
3.394 | -34.0924 | -6.296
3.4001 | -34.0925 | -6.3068
3.4062 | -34.0925 | -6.3171
3.4123 | -34.0926 | -6.327
3.4184 | -34.0926 | -6.3381
3.4245 | -34.0928 | -6.3486
3.4306 | -34.0928 | -6.3585
3.4367 | -34.0929 | -6.3684
3.4428 | -34.0932 | -6.3777
3.4489 | -34.0932 | -6.3888
3.455 | -34.0931 | -6.3993
3.4611 | -34.0932 | -6.4104
3.4672 | -34.0933 | -6.4216
3.4733 | -34.0934 | -6.4319
3.4794 | -34.0934 | -6.4425
3.4855 | -34.0936 | -6.4522
3.4916 | -34.0938 | -6.462
3.4977 | -34.0938 | -6.473
3.5038 | -34.0939 | -6.4831
3.5099 | -34.094 | -6.4928
3.5161 | -34.0942 | -6.5024
3.5222 | -34.0943 | -6.5125
3.5283 | -34.0943 | -6.5226
3.5344 | -34.0943 | -6.5334
3.5405 | -34.0945 | -6.5434
3.5466 | -34.0946 | -6.5534
3.5527 | -34.0947 | -6.5641
3.5588 | -34.0949 | -6.574
3.5649 | -34.0949 | -6.5842
3.571 | -34.095 | -6.5945
3.5771 | -34.0951 | -6.6053
3.5832 | -34.0953 | -6.6149
3.5893 | -34.0955 | -6.6249
3.5954 | -34.0956 | -6.6354
3.6015 | -34.0958 | -6.6458
3.6076 | -34.096 | -6.6556
3.6137 | -34.0961 | -6.6655
3.6198 | -34.0965 | -6.6747
3.6259 | -34.0966 | -6.6855
3.632 | -34.0968 | -6.6961
3.6381 | -34.097 | -6.7061
3.6442 | -34.0972 | -6.7166
3.6503 | -34.0975 | -6.7258
3.6565 | -34.0978 | -6.7374
3.6626 | -34.0979 | -6.7481
3.6687 | -34.0982 | -6.7587
3.6748 | -34.0985 | -6.7687
3.6809 | -34.0988 | -6.7792
3.687 | -34.099 | -6.7903
3.6931 | -34.0994 | -6.8003
3.6992 | -34.0997 | -6.811
3.7053 | -34.0999 | -6.8211
3.7114 | -34.1001 | -6.8323
3.7175 | -34.1004 | -6.8423
3.7236 | -34.1006 | -6.8513
3.7297 | -34.1008 | -6.8621
3.7358 | -34.1011 | -6.8736
3.7419 | -34.1013 | -6.8851
3.748 | -34.1014 | -6.8957
3.7541 | -34.1016 | -6.9066
3.7602 | -34.102 | -6.9174
3.7663 | -34.1022 | -6.9276
3.7724 | -34.1025 | -6.9385
3.7785 | -34.1029 | -6.9489
3.7846 | -34.1031 | -6.9595
3.7907 | -34.1035 | -6.9705
3.7969 | -34.1037 | -6.9818
3.803 | -34.1041 | -6.9919
3.8091 | -34.1044 | -7.0023
3.8152 | -34.1047 | -7.0133
3.8213 | -34.1049 | -7.0248
3.8274 | -34.1053 | -7.0352
3.8335 | -34.1057 | -7.0466
3.8396 | -34.1061 | -7.058
3.8457 | -34.1064 | -7.0683
3.8518 | -34.1069 | -7.0791
3.8579 | -34.1073 | -7.0897
3.864 | -34.1077 | -7.1004
3.8701 | -34.1081 | -7.1114
3.8762 | -34.1084 | -7.1224
3.8823 | -34.1087 | -7.133
3.8884 | -34.1091 | -7.1434
3.8945 | -34.1095 | -7.1546
3.9006 | -34.1099 | -7.1661
3.9067 | -34.1103 | -7.1778
3.9128 | -34.1106 | -7.1889
3.9189 | -34.111 | -7.1995
3.925 | -34.1113 | -7.2104
3.9311 | -34.1117 | -7.2204
3.9372 | -34.1121 | -7.2318
3.9434 | -34.1125 | -7.2423
3.9495 | -34.1128 | -7.2537
3.9556 | -34.1133 | -7.2648
3.9617 | -34.1136 | -7.2755
3.9678 | -34.114 | -7.2862
3.9739 | -34.1144 | -7.2972
3.98 | -34.1148 | -7.3088
3.9861 | -34.1151 | -7.3193
3.9922 | -34.1153 | -7.3308
3.9983 | -34.1157 | -7.3408
4.0044 | -34.116 | -7.3506
4.0105 | -34.1164 | -7.3614
4.0166 | -34.1169 | -7.3724
4.0227 | -34.1173 | -7.3828
4.0288 | -34.1175 | -7.3941
4.0349 | -34.1178 | -7.404
4.041 | -34.1182 | -7.4149
4.0471 | -34.1184 | -7.4252
4.0532 | -34.1189 | -7.4362
4.0593 | -34.1192 | -7.4476
4.0654 | -34.1195 | -7.4583
4.0715 | -34.1198 | -7.4695
4.0776 | -34.1202 | -7.4799
4.0838 | -34.1205 | -7.4912
4.0899 | -34.1209 | -7.5023
4.096 | -34.1212 | -7.5127
4.1021 | -34.1215 | -7.5234
4.1082 | -34.1219 | -7.5336
4.1143 | -34.1222 | -7.5436
4.1204 | -34.1226 | -7.5542
4.1265 | -34.123 | -7.5646
4.1326 | -34.1234 | -7.5749
4.1387 | -34.1238 | -7.5859
4.1448 | -34.1241 | -7.5968
4.1509 | -34.1244 | -7.6074
4.157 | -34.1246 | -7.6186
4.1631 | -34.125 | -7.6297
4.1692 | -34.1253 | -7.6403
4.1753 | -34.1257 | -7.6511
4.1814 | -34.1259 | -7.6609
4.1875 | -34.1262 | -7.6721
4.1936 | -34.1265 | -7.6833
4.1997 | -34.1269 | -7.6934
4.2058 | -34.1271 | -7.7048
4.2119 | -34.1274 | -7.7146
4.218 | -34.1277 | -7.7247
4.2241 | -34.128 | -7.7344
4.2303 | -34.1284 | -7.7454
4.2364 | -34.1288 | -7.7567
4.2425 | -34.129 | -7.7671
4.2486 | -34.1293 | -7.7782
4.2547 | -34.1294 | -7.7875
4.2608 | -34.1297 | -7.7975
4.2669 | -34.13 | -7.8078
4.273 | -34.1303 | -7.8185
4.2791 | -34.1305 | -7.8292
4.2852 | -34.1307 | -7.8395
4.2913 | -34.1309 | -7.8488
4.2974 | -34.1311 | -7.8597
4.3035 | -34.1314 | -7.8705
4.3096 | -34.1315 | -7.8809
4.3157 | -34.1318 | -7.8909
4.3218 | -34.132 | -7.9008
4.3279 | -34.1323 | -7.9112
4.334 | -34.1325 | -7.9211
4.3401 | -34.1327 | -7.931
4.3462 | -34.133 | -7.9417
4.3523 | -34.1334 | -7.9514
4.3584 | -34.1335 | -7.9601
4.3645 | -34.1337 | -7.9698
4.3707 | -34.1339 | -7.9792
4.3768 | -34.1342 | -7.9895
4.3829 | -34.1343 | -7.9999
4.389 | -34.1346 | -8.0096
4.3951 | -34.1349 | -8.0193
4.4012 | -34.1351 | -8.0292
4.4073 | -34.1353 | -8.0393
4.4134 | -34.1355 | -8.0494
4.4195 | -34.1358 | -8.0591
4.4256 | -34.1361 | -8.0687
4.4317 | -34.1363 | -8.0783
4.4378 | -34.1365 | -8.0879
4.4439 | -34.1367 | -8.0974
4.45 | -34.1368 | -8.1074
4.4561 | -34.1369 | -8.1173
4.4622 | -34.1371 | -8.1269
4.4683 | -34.1373 | -8.1363
4.4744 | -34.1375 | -8.1465
4.4805 | -34.1378 | -8.1557
4.4866 | -34.138 | -8.1651
4.4927 | -34.1382 | -8.1753
4.4988 | -34.1384 | -8.1848
4.5049 | -34.1387 | -8.194
4.511 | -34.1389 | -8.2032
4.5172 | -34.1392 | -8.2131
4.5233 | -34.1394 | -8.2236
4.5294 | -34.1397 | -8.2331
4.5355 | -34.1399 | -8.2426
4.5416 | -34.1402 | -8.2518
4.5477 | -34.1404 | -8.2612
4.5538 | -34.1406 | -8.2701
4.5599 | -34.1409 | -8.2797
4.566 | -34.1411 | -8.2895
4.5721 | -34.1413 | -8.2988
4.5782 | -34.1416 | -8.3087
4.5843 | -34.1418 | -8.3184
4.5904 | -34.1422 | -8.3279
4.5965 | -34.1425 | -8.3368
4.6026 | -34.1428 | -8.3474
4.6087 | -34.143 | -8.3565
4.6148 | -34.1432 | -8.3671
4.6209 | -34.1436 | -8.3758
4.627 | -34.1439 | -8.3859
4.6331 | -34.144 | -8.395
4.6392 | -34.1443 | -8.4044
4.6453 | -34.1445 | -8.4139
4.6514 | -34.1448 | -8.4235
4.6576 | -34.1449 | -8.434
4.6637 | -34.1451 | -8.442
4.6698 | -34.1453 | -8.4508
4.6759 | -34.1455 | -8.4607
4.682 | -34.1457 | -8.47
4.6881 | -34.1459 | -8.479
4.6942 | -34.1462 | -8.4884
4.7003 | -34.1464 | -8.4979
4.7064 | -34.1466 | -8.5073
4.7125 | -34.147 | -8.5169
4.7186 | -34.1471 | -8.5256
4.7247 | -34.1475 | -8.5353
4.7308 | -34.1478 | -8.5448
4.7369 | -34.148 | -8.5545
4.743 | -34.1482 | -8.5648
4.7491 | -34.1484 | -8.5746
4.7552 | -34.1486 | -8.584
4.7613 | -34.1488 | -8.5937
4.7674 | -34.1491 | -8.6025
4.7735 | -34.1492 | -8.6127
4.7796 | -34.1495 | -8.6226
4.7857 | -34.1496 | -8.6324
4.7918 | -34.1499 | -8.6413
4.7979 | -34.1502 | -8.6505
4.8041 | -34.1504 | -8.6599
4.8102 | -34.1507 | -8.6691
4.8163 | -34.1509 | -8.6787
4.8224 | -34.1511 | -8.6876
4.8285 | -34.1514 | -8.6968
4.8346 | -34.1517 | -8.7064
4.8407 | -34.1519 | -8.7159
4.8468 | -34.1522 | -8.7257
4.8529 | -34.1524 | -8.7356
4.859 | -34.1526 | -8.7453
4.8651 | -34.1529 | -8.7548
4.8712 | -34.1531 | -8.7643
4.8773 | -34.1533 | -8.774
4.8834 | -34.1536 | -8.7838
4.8895 | -34.1539 | -8.7938
4.8956 | -34.1541 | -8.8038
4.9017 | -34.1543 | -8.8134
4.9078 | -34.1545 | -8.8234
4.9139 | -34.1548 | -8.8322
4.92 | -34.155 | -8.8414
4.9261 | -34.1554 | -8.8503
4.9322 | -34.1557 | -8.8597
4.9383 | -34.1561 | -8.8689
4.9445 | -34.1564 | -8.8791
4.9506 | -34.1566 | -8.8881
4.9567 | -34.1569 | -8.8975
4.9628 | -34.1571 | -8.907
4.9689 | -34.1573 | -8.917
4.975 | -34.1575 | -8.9266
4.9811 | -34.1578 | -8.9364
4.9872 | -34.1581 | -8.9462
4.9933 | -34.1585 | -8.956
4.9994 | -34.1588 | -8.966
5.0055 | -34.1592 | -8.9758
5.0116 | -34.1595 | -8.9861
5.0177 | -34.1599 | -8.9962
5.0238 | -34.1602 | -9.0067
5.0299 | -34.1607 | -9.0177
5.036 | -34.1609 | -9.0273
5.0421 | -34.1612 | -9.0371
5.0482 | -34.1615 | -9.0474
5.0543 | -34.1618 | -9.0568
5.0604 | -34.1622 | -9.0671
5.0665 | -34.1625 | -9.0768
5.0726 | -34.1628 | -9.0872
5.0787 | -34.1631 | -9.0975
5.0848 | -34.1635 | -9.1076
5.091 | -34.1638 | -9.1175
5.0971 | -34.1641 | -9.1276
5.1032 | -34.1645 | -9.138
5.1093 | -34.1648 | -9.148
5.1154 | -34.165 | -9.158
5.1215 | -34.1653 | -9.1681
5.1276 | -34.1656 | -9.1779
5.1337 | -34.1661 | -9.1877
5.1398 | -34.1663 | -9.1975
5.1459 | -34.1667 | -9.208
5.152 | -34.167 | -9.2189
5.1581 | -34.1673 | -9.2296
5.1642 | -34.1676 | -9.2396
5.1703 | -34.1679 | -9.2496
5.1764 | -34.1683 | -9.2596
5.1825 | -34.1686 | -9.2699
5.1886 | -34.1689 | -9.2797
5.1947 | -34.1692 | -9.2904
5.2008 | -34.1696 | -9.3003
5.2069 | -34.1699 | -9.3114
5.213 | -34.1703 | -9.321
5.2191 | -34.1707 | -9.3314
5.2252 | -34.1709 | -9.3417
5.2314 | -34.1712 | -9.3524
5.2375 | -34.1716 | -9.3623
5.2436 | -34.1719 | -9.3725
5.2497 | -34.1723 | -9.3824
5.2558 | -34.1726 | -9.3933
5.2619 | -34.1729 | -9.4041
5.268 | -34.1733 | -9.4146
5.2741 | -34.1737 | -9.4256
5.2802 | -34.174 | -9.4364
5.2863 | -34.1744 | -9.4466
5.2924 | -34.1747 | -9.4577
5.2985 | -34.1751 | -9.4689
5.3046 | -34.1755 | -9.4799
5.3107 | -34.1759 | -9.4907
5.3168 | -34.1764 | -9.5011
5.3229 | -34.1767 | -9.5111
5.329 | -34.177 | -9.5211
5.3351 | -34.1775 | -9.5313
5.3412 | -34.1779 | -9.5414
5.3473 | -34.1782 | -9.5522
5.3534 | -34.1786 | -9.563
5.3595 | -34.1789 | -9.5737
5.3656 | -34.1793 | -9.5845
5.3717 | -34.1796 | -9.5951
5.3779 | -34.18 | -9.6059
5.384 | -34.1803 | -9.6166
5.3901 | -34.1807 | -9.6273
5.3962 | -34.181 | -9.638
5.4023 | -34.1814 | -9.6487
5.4084 | -34.1818 | -9.6591
5.4145 | -34.1821 | -9.6699
5.4206 | -34.1824 | -9.681
5.4267 | -34.1828 | -9.6922
5.4328 | -34.1832 | -9.7024
5.4389 | -34.1835 | -9.7126
5.445 | -34.1838 | -9.7239
5.4511 | -34.1842 | -9.7342
5.4572 | -34.1846 | -9.7453
5.4633 | -34.185 | -9.7557
5.4694 | -34.1852 | -9.7661
5.4755 | -34.1856 | -9.7768
5.4816 | -34.186 | -9.7874
5.4877 | -34.1863 | -9.798
5.4938 | -34.1867 | -9.8084
5.4999 | -34.187 | -9.819
5.506 | -34.1874 | -9.8296
5.5121 | -34.1877 | -9.8405
5.5183 | -34.1881 | -9.8506
5.5244 | -34.1883 | -9.8614
5.5305 | -34.1887 | -9.8722
5.5366 | -34.1891 | -9.8825
5.5427 | -34.1894 | -9.8937
5.5488 | -34.1897 | -9.9046
5.5549 | -34.19 | -9.9152
5.561 | -34.1904 | -9.9263
5.5671 | -34.1908 | -9.9367
5.5732 | -34.1911 | -9.9475
5.5793 | -34.1914 | -9.9589
5.5854 | -34.1919 | -9.9697
5.5915 | -34.1921 | -9.9806
5.5976 | -34.1924 | -9.9915
5.6037 | -34.1928 | -10.0021
5.6098 | -34.1931 | -10.0126
5.6159 | -34.1934 | -10.0234
5.622 | -34.1938 | -10.0338
5.6281 | -34.1941 | -10.0446
5.6342 | -34.1945 | -10.0554
5.6403 | -34.1948 | -10.0657
5.6464 | -34.1952 | -10.0764
5.6525 | -34.1954 | -10.0869
5.6586 | -34.1959 | -10.0969
5.6648 | -34.1962 | -10.108
5.6709 | -34.1965 | -10.119
5.677 | -34.1969 | -10.1303
5.6831 | -34.1972 | -10.1406
5.6892 | -34.1975 | -10.1513
5.6953 | -34.1979 | -10.1619
5.7014 | -34.1982 | -10.173
5.7075 | -34.1985 | -10.1839
5.7136 | -34.1989 | -10.1943
5.7197 | -34.1992 | -10.2049
5.7258 | -34.1994 | -10.2153
5.7319 | -34.1998 | -10.226
5.738 | -34.2002 | -10.2367
5.7441 | -34.2005 | -10.2473
5.7502 | -34.2009 | -10.2575
5.7563 | -34.2012 | -10.268
5.7624 | -34.2016 | -10.2793
5.7685 | -34.2018 | -10.2898
5.7746 | -34.2021 | -10.3003
5.7807 | -34.2024 | -10.3107
5.7868 | -34.2027 | -10.3213
5.7929 | -34.203 | -10.3318
5.799 | -34.2033 | -10.3423
5.8052 | -34.2036 | -10.353
5.8113 | -34.2039 | -10.3636
5.8174 | -34.2042 | -10.3743
5.8235 | -34.2045 | -10.3847
5.8296 | -34.2048 | -10.3958
5.8357 | -34.2051 | -10.4064
5.8418 | -34.2054 | -10.4172
5.8479 | -34.2056 | -10.428
5.854 | -34.206 | -10.4389
5.8601 | -34.2063 | -10.4493
5.8662 | -34.2066 | -10.4602
5.8723 | -34.2069 | -10.4702
5.8784 | -34.2072 | -10.4808
5.8845 | -34.2075 | -10.4915
5.8906 | -34.2078 | -10.502
5.8967 | -34.2081 | -10.5122
5.9028 | -34.2085 | -10.5228
5.9089 | -34.2088 | -10.5328
5.915 | -34.2091 | -10.5434
5.9211 | -34.2094 | -10.5539
5.9272 | -34.2097 | -10.5644
5.9333 | -34.21 | -10.5743
5.9394 | -34.2104 | -10.5847
5.9455 | -34.2107 | -10.5948
5.9517 | -34.2111 | -10.6053
5.9578 | -34.2114 | -10.6161
5.9639 | -34.2118 | -10.6263
5.97 | -34.2121 | -10.6369
5.9761 | -34.2124 | -10.6471
5.9822 | -34.2128 | -10.6577
5.9883 | -34.2131 | -10.668
5.9944 | -34.2134 | -10.6785
6.0005 | -34.2138 | -10.6893
6.0066 | -34.2141 | -10.6994
6.0127 | -34.2144 | -10.7092
6.0188 | -34.2148 | -10.72
6.0249 | -34.2151 | -10.7304
6.031 | -34.2154 | -10.7406
6.0371 | -34.2158 | -10.7513
6.0432 | -34.2161 | -10.7617
6.0493 | -34.2165 | -10.772
6.0554 | -34.2169 | -10.7823
6.0615 | -34.2172 | -10.7931
6.0676 | -34.2175 | -10.8041
6.0737 | -34.2179 | -10.8147
6.0798 | -34.2181 | -10.825
6.0859 | -34.2185 | -10.8354
6.0921 | -34.2189 | -10.8456
6.0982 | -34.2192 | -10.8562
6.1043 | -34.2196 | -10.8665
6.1104 | -34.22 | -10.8773
6.1165 | -34.2203 | -10.887
6.1226 | -34.2207 | -10.897
6.1287 | -34.2211 | -10.9074
6.1348 | -34.2215 | -10.9179
6.1409 | -34.2218 | -10.9282
6.147 | -34.2221 | -10.9384
6.1531 | -34.2225 | -10.9488
6.1592 | -34.2228 | -10.9591
6.1653 | -34.2232 | -10.9693
6.1714 | -34.2235 | -10.9797
6.1775 | -34.2239 | -10.9903
6.1836 | -34.2242 | -11.0005
6.1897 | -34.2245 | -11.0109
6.1958 | -34.2248 | -11.0214
6.2019 | -34.2252 | -11.032
6.208 | -34.2256 | -11.0427
6.2141 | -34.226 | -11.0533
6.2202 | -34.2263 | -11.064
6.2263 | -34.2267 | -11.0743
6.2325 | -34.2271 | -11.0842
6.2386 | -34.2275 | -11.0942
6.2447 | -34.2278 | -11.1048
6.2508 | -34.2281 | -11.115
6.2569 | -34.2285 | -11.1256
6.263 | -34.2289 | -11.1357
6.2691 | -34.2292 | -11.1459
6.2752 | -34.2296 | -11.1562
6.2813 | -34.2299 | -11.1665
6.2874 | -34.2303 | -11.1768
6.2935 | -34.2307 | -11.187
6.2996 | -34.231 | -11.1974
6.3057 | -34.2314 | -11.2076
6.3118 | -34.2317 | -11.2179
6.3179 | -34.2321 | -11.2282
6.324 | -34.2325 | -11.2388
6.3301 | -34.2329 | -11.2489
6.3362 | -34.2332 | -11.2593
6.3423 | -34.2337 | -11.2697
6.3484 | -34.2341 | -11.2802
6.3545 | -34.2344 | -11.29
6.3606 | -34.2348 | -11.3
6.3667 | -34.2352 | -11.3101
6.3728 | -34.2356 | -11.3204
6.379 | -34.2359 | -11.3307
6.3851 | -34.2363 | -11.3409
6.3912 | -34.2367 | -11.351
6.3973 | -34.2371 | -11.3611
6.4034 | -34.2375 | -11.3713
6.4095 | -34.2378 | -11.3817
6.4156 | -34.2382 | -11.3917
6.4217 | -34.2385 | -11.402
6.4278 | -34.2389 | -11.4122
6.4339 | -34.2393 | -11.4224
6.44 | -34.2397 | -11.4323
6.4461 | -34.24 | -11.443
6.4522 | -34.2404 | -11.4527
6.4583 | -34.2407 | -11.463
6.4644 | -34.241 | -11.4736
6.4705 | -34.2415 | -11.4839
6.4766 | -34.2419 | -11.4945
6.4827 | -34.2422 | -11.5051
6.4888 | -34.2425 | -11.5152
6.4949 | -34.243 | -11.525
6.501 | -34.2433 | -11.5356
6.5071 | -34.2437 | -11.5459
6.5132 | -34.2441 | -11.556
6.5194 | -34.2445 | -11.5662
6.5255 | -34.2449 | -11.5761
6.5316 | -34.2452 | -11.5865
6.5377 | -34.2456 | -11.597
6.5438 | -34.246 | -11.6071
6.5499 | -34.2464 | -11.6171
6.556 | -34.2467 | -11.6272
6.5621 | -34.2471 | -11.6376
6.5682 | -34.2474 | -11.6477
6.5743 | -34.2478 | -11.6586
6.5804 | -34.2481 | -11.6689
6.5865 | -34.2484 | -11.6793
6.5926 | -34.2489 | -11.6895
6.5987 | -34.2492 | -11.6994
6.6048 | -34.2496 | -11.7102
6.6109 | -34.25 | -11.7203
6.617 | -34.2503 | -11.7313
6.6231 | -34.2506 | -11.7417
6.6292 | -34.2511 | -11.7519
6.6353 | -34.2514 | -11.763
6.6414 | -34.2519 | -11.7735
6.6475 | -34.2523 | -11.7839
6.6536 | -34.2526 | -11.7939
6.6597 | -34.2531 | -11.8039
6.6659 | -34.2534 | -11.814
6.672 | -34.2538 | -11.8243
6.6781 | -34.2542 | -11.8344
6.6842 | -34.2546 | -11.8447
6.6903 | -34.255 | -11.8549
6.6964 | -34.2554 | -11.8652
6.7025 | -34.2557 | -11.8754
6.7086 | -34.2561 | -11.8854
6.7147 | -34.2564 | -11.8958
6.7208 | -34.2568 | -11.9065
6.7269 | -34.2571 | -11.9165
6.733 | -34.2574 | -11.926
6.7391 | -34.2579 | -11.9364
6.7452 | -34.2583 | -11.9462
6.7513 | -34.2587 | -11.9561
6.7574 | -34.2591 | -11.9665
6.7635 | -34.2595 | -11.9763
6.7696 | -34.26 | -11.9866
6.7757 | -34.2604 | -11.9971
6.7818 | -34.2607 | -12.0073
6.7879 | -34.2611 | -12.0176
6.794 | -34.2615 | -12.0276
6.8001 | -34.2619 | -12.0377
6.8063 | -34.2623 | -12.0481
6.8124 | -34.2626 | -12.0584
6.8185 | -34.263 | -12.0686
6.8246 | -34.2633 | -12.0788
6.8307 | -34.2637 | -12.0889
6.8368 | -34.2641 | -12.0994
6.8429 | -34.2645 | -12.1095
6.849 | -34.2649 | -12.1198
6.8551 | -34.2653 | -12.13
6.8612 | -34.2657 | -12.1401
6.8673 | -34.266 | -12.1501
6.8734 | -34.2664 | -12.1609
6.8795 | -34.2668 | -12.1713
6.8856 | -34.2672 | -12.1811
6.8917 | -34.2676 | -12.1908
6.8978 | -34.2679 | -12.2009
6.9039 | -34.2683 | -12.2109
6.91 | -34.2687 | -12.2212
6.9161 | -34.2691 | -12.2312
6.9222 | -34.2695 | -12.2414
6.9283 | -34.2698 | -12.2512
6.9344 | -34.2702 | -12.2612
6.9405 | -34.2706 | -12.2712
6.9466 | -34.271 | -12.2816
6.9528 | -34.2713 | -12.2918
6.9589 | -34.2717 | -12.3019
6.965 | -34.2721 | -12.3121
6.9711 | -34.2725 | -12.3223
6.9772 | -34.273 | -12.3322
6.9833 | -34.2733 | -12.3425
6.9894 | -34.2736 | -12.3531
6.9955 | -34.274 | -12.3631
7.0016 | -34.2744 | -12.3736
7.0077 | -34.2748 | -12.384
7.0138 | -34.2752 | -12.3944
7.0199 | -34.2755 | -12.4044
7.026 | -34.2758 | -12.4144
7.0321 | -34.2762 | -12.4245
7.0382 | -34.2765 | -12.4344
7.0443 | -34.277 | -12.4448
7.0504 | -34.2773 | -12.4553
7.0565 | -34.2777 | -12.4656
7.0626 | -34.278 | -12.4755
7.0687 | -34.2784 | -12.4852
7.0748 | -34.2788 | -12.4952
7.0809 | -34.2792 | -12.5054
7.087 | -34.2795 | -12.5156
7.0932 | -34.2799 | -12.5257
7.0993 | -34.2802 | -12.5358
7.1054 | -34.2806 | -12.5457
7.1115 | -34.2809 | -12.5557
7.1176 | -34.2812 | -12.5656
7.1237 | -34.2815 | -12.5757
7.1298 | -34.2819 | -12.5855
7.1359 | -34.2822 | -12.5953
7.142 | -34.2825 | -12.6054
7.1481 | -34.2829 | -12.6152
7.1542 | -34.2832 | -12.6252
7.1603 | -34.2836 | -12.6351
7.1664 | -34.284 | -12.6452
7.1725 | -34.2844 | -12.6551
7.1786 | -34.2848 | -12.6652
7.1847 | -34.2851 | -12.6752
7.1908 | -34.2854 | -12.6853
7.1969 | -34.2858 | -12.6954
7.203 | -34.2861 | -12.7054
7.2091 | -34.2865 | -12.7152
7.2152 | -34.2868 | -12.7252
7.2213 | -34.2872 | -12.735
7.2274 | -34.2875 | -12.745
7.2335 | -34.2879 | -12.755
7.2397 | -34.2883 | -12.7649
7.2458 | -34.2886 | -12.7745
7.2519 | -34.289 | -12.7845
7.258 | -34.2893 | -12.7947
7.2641 | -34.2897 | -12.8045
7.2702 | -34.29 | -12.8148
7.2763 | -34.2904 | -12.8246
7.2824 | -34.2908 | -12.8345
7.2885 | -34.2912 | -12.8444
7.2946 | -34.2916 | -12.8545
7.3007 | -34.2918 | -12.8646
7.3068 | -34.2922 | -12.8748
7.3129 | -34.2925 | -12.8846
7.319 | -34.2929 | -12.8946
7.3251 | -34.2933 | -12.9043
7.3312 | -34.2936 | -12.9139
7.3373 | -34.294 | -12.9236
7.3434 | -34.2945 | -12.9333
7.3495 | -34.2948 | -12.9428
7.3556 | -34.2952 | -12.9531
7.3617 | -34.2955 | -12.9635
7.3678 | -34.2959 | -12.9735
7.3739 | -34.2962 | -12.9831
7.3801 | -34.2965 | -12.993
7.3862 | -34.2969 | -13.0029
7.3923 | -34.2973 | -13.0127
7.3984 | -34.2976 | -13.0225
7.4045 | -34.298 | -13.0323
7.4106 | -34.2984 | -13.0421
7.4167 | -34.2987 | -13.0522
7.4228 | -34.2991 | -13.062
7.4289 | -34.2994 | -13.0717
7.435 | -34.2998 | -13.0813
7.4411 | -34.3002 | -13.0909
7.4472 | -34.3005 | -13.1009
7.4533 | -34.3009 | -13.1107
7.4594 | -34.3013 | -13.1209
7.4655 | -34.3017 | -13.1306
7.4716 | -34.302 | -13.1407
7.4777 | -34.3023 | -13.1501
7.4838 | -34.3027 | -13.1601
7.4899 | -34.3031 | -13.1699
7.496 | -34.3035 | -13.1797
7.5021 | -34.3038 | -13.1895
7.5082 | -34.3042 | -13.1994
7.5143 | -34.3045 | -13.2092
7.5204 | -34.3049 | -13.2188
7.5266 | -34.3052 | -13.2287
7.5327 | -34.3056 | -13.2383
7.5388 | -34.3059 | -13.2482
7.5449 | -34.3062 | -13.2582
7.551 | -34.3066 | -13.2682
7.5571 | -34.3069 | -13.278
7.5632 | -34.3073 | -13.2877
7.5693 | -34.3077 | -13.2977
7.5754 | -34.3081 | -13.3075
7.5815 | -34.3084 | -13.317
7.5876 | -34.3088 | -13.3268
7.5937 | -34.3091 | -13.3366
7.5998 | -34.3095 | -13.3465
7.6059 | -34.3098 | -13.3563
7.612 | -34.3102 | -13.3661
7.6181 | -34.3105 | -13.3758
7.6242 | -34.3109 | -13.3857
7.6303 | -34.3113 | -13.3953
7.6364 | -34.3116 | -13.4053
7.6425 | -34.312 | -13.4154
7.6486 | -34.3123 | -13.4253
7.6547 | -34.3127 | -13.4351
7.6608 | -34.3131 | -13.4449
7.667 | -34.3134 | -13.4542
7.6731 | -34.3137 | -13.4641
7.6792 | -34.3141 | -13.4741
7.6853 | -34.3144 | -13.4838
7.6914 | -34.3147 | -13.4937
7.6975 | -34.315 | -13.5037
7.7036 | -34.3154 | -13.5136
7.7097 | -34.3158 | -13.5234
7.7158 | -34.3161 | -13.5331
7.7219 | -34.3164 | -13.5429
7.728 | -34.3168 | -13.5528
7.7341 | -34.3171 | -13.5627
7.7402 | -34.3175 | -13.5724
7.7463 | -34.3178 | -13.5822
7.7524 | -34.3181 | -13.5923
7.7585 | -34.3184 | -13.6022
7.7646 | -34.3188 | -13.6118
7.7707 | -34.3191 | -13.6213
7.7768 | -34.3193 | -13.631
7.7829 | -34.3197 | -13.6409
7.789 | -34.32 | -13.6505
7.7951 | -34.3203 | -13.6607
7.8012 | -34.3207 | -13.6708
7.8073 | -34.321 | -13.6805
7.8135 | -34.3213 | -13.6904
7.8196 | -34.3217 | -13.7002
7.8257 | -34.3221 | -13.7104
7.8318 | -34.3224 | -13.7205
7.8379 | -34.3227 | -13.73
7.844 | -34.323 | -13.7394
7.8501 | -34.3233 | -13.7494
7.8562 | -34.3236 | -13.7598
7.8623 | -34.324 | -13.7697
7.8684 | -34.3243 | -13.7792
7.8745 | -34.3246 | -13.789
7.8806 | -34.3249 | -13.7984
7.8867 | -34.3253 | -13.808
7.8928 | -34.3257 | -13.8176
7.8989 | -34.326 | -13.827
7.905 | -34.3264 | -13.8372
7.9111 | -34.3267 | -13.847
7.9172 | -34.3271 | -13.8566
7.9233 | -34.3273 | -13.8665
7.9294 | -34.3277 | -13.8764
7.9355 | -34.328 | -13.8858
7.9416 | -34.3283 | -13.8955
7.9477 | -34.3287 | -13.9049
7.9539 | -34.329 | -13.9144
7.96 | -34.3293 | -13.9244
7.9661 | -34.3297 | -13.9342
7.9722 | -34.33 | -13.9441
7.9783 | -34.3303 | -13.9539
7.9844 | -34.3307 | -13.964
7.9905 | -34.331 | -13.9738
7.9966 | -34.3313 | -13.9836
8.0027 | -34.3316 | -13.9936
8.0088 | -34.332 | -14.0031
8.0149 | -34.3323 | -14.0126
8.021 | -34.3326 | -14.0224
8.0271 | -34.333 | -14.0326
8.0332 | -34.3333 | -14.0423
8.0393 | -34.3337 | -14.0522
8.0454 | -34.3339 | -14.0622
8.0515 | -34.3343 | -14.072
8.0576 | -34.3346 | -14.082
8.0637 | -34.3349 | -14.0914
8.0698 | -34.3353 | -14.1008
8.0759 | -34.3356 | -14.1103
8.082 | -34.3359 | -14.1202
8.0881 | -34.3362 | -14.1298
8.0942 | -34.3366 | -14.1399
8.1004 | -34.3369 | -14.15
8.1065 | -34.3372 | -14.1597
8.1126 | -34.3376 | -14.1694
8.1187 | -34.3379 | -14.1793
8.1248 | -34.3383 | -14.1891
8.1309 | -34.3386 | -14.1989
8.137 | -34.3389 | -14.2085
8.1431 | -34.3393 | -14.2184
8.1492 | -34.3396 | -14.2283
8.1553 | -34.3399 | -14.2381
8.1614 | -34.3403 | -14.248
8.1675 | -34.3406 | -14.2577
8.1736 | -34.3409 | -14.2678
8.1797 | -34.3412 | -14.2778
8.1858 | -34.3415 | -14.2876
8.1919 | -34.3418 | -14.2972
8.198 | -34.3421 | -14.3067
8.2041 | -34.3425 | -14.3167
8.2102 | -34.3429 | -14.3264
8.2163 | -34.3432 | -14.3361
8.2224 | -34.3435 | -14.3461
8.2285 | -34.3439 | -14.3555
8.2346 | -34.3442 | -14.3652
8.2408 | -34.3445 | -14.375
8.2469 | -34.3449 | -14.3846
8.253 | -34.3452 | -14.3944
8.2591 | -34.3455 | -14.4039
8.2652 | -34.3459 | -14.4136
8.2713 | -34.3463 | -14.4234
8.2774 | -34.3466 | -14.4332
8.2835 | -34.3469 | -14.4434
8.2896 | -34.3473 | -14.4527
8.2957 | -34.3476 | -14.4625
8.3018 | -34.3479 | -14.4722
8.3079 | -34.3483 | -14.4819
8.314 | -34.3487 | -14.4914
8.3201 | -34.349 | -14.5012
8.3262 | -34.3493 | -14.5109
8.3323 | -34.3497 | -14.5206
8.3384 | -34.35 | -14.5307
8.3445 | -34.3504 | -14.5401
8.3506 | -34.3507 | -14.5501
8.3567 | -34.351 | -14.56
8.3628 | -34.3514 | -14.5697
8.3689 | -34.3517 | -14.5795
8.375 | -34.352 | -14.5893
8.3812 | -34.3524 | -14.5989
8.3873 | -34.3527 | -14.6088
8.3934 | -34.353 | -14.6186
8.3995 | -34.3534 | -14.6284
8.4056 | -34.3537 | -14.6381
8.4117 | -34.354 | -14.648
8.4178 | -34.3544 | -14.6574
8.4239 | -34.3547 | -14.667
8.43 | -34.3551 | -14.6767
8.4361 | -34.3554 | -14.6868
8.4422 | -34.3557 | -14.6964
8.4483 | -34.356 | -14.706
8.4544 | -34.3564 | -14.7161
8.4605 | -34.3567 | -14.7257
8.4666 | -34.357 | -14.7355
8.4727 | -34.3573 | -14.7452
8.4788 | -34.3577 | -14.7548
8.4849 | -34.358 | -14.7648
8.491 | -34.3583 | -14.7744
8.4971 | -34.3587 | -14.7842
8.5032 | -34.359 | -14.794
8.5093 | -34.3593 | -14.8037
8.5154 | -34.3597 | -14.8133
8.5215 | -34.36 | -14.8229
8.5277 | -34.3603 | -14.8328
8.5338 | -34.3606 | -14.8423
8.5399 | -34.361 | -14.8523
8.546 | -34.3614 | -14.8617
8.5521 | -34.3617 | -14.8716
8.5582 | -34.362 | -14.8814
8.5643 | -34.3623 | -14.8909
8.5704 | -34.3627 | -14.9004
8.5765 | -34.363 | -14.9101
8.5826 | -34.3635 | -14.9198
8.5887 | -34.3638 | -14.9294
8.5948 | -34.3641 | -14.9389
8.6009 | -34.3645 | -14.9485
8.607 | -34.3648 | -14.9584
8.6131 | -34.3651 | -14.9683
8.6192 | -34.3654 | -14.9777
8.6253 | -34.3658 | -14.9875
8.6314 | -34.3662 | -14.9972
8.6375 | -34.3665 | -15.0071
8.6436 | -34.3669 | -15.0168
8.6497 | -34.3672 | -15.0265
8.6558 | -34.3675 | -15.036
8.6619 | -34.3678 | -15.0459
8.6681 | -34.3682 | -15.0555
8.6742 | -34.3685 | -15.0653
8.6803 | -34.3689 | -15.0753
8.6864 | -34.3693 | -15.0851
8.6925 | -34.3696 | -15.0948
8.6986 | -34.37 | -15.1048
8.7047 | -34.3703 | -15.1147
8.7108 | -34.3706 | -15.1245
8.7169 | -34.371 | -15.134
8.723 | -34.3714 | -15.1437
8.7291 | -34.3717 | -15.1533
8.7352 | -34.3721 | -15.1633
8.7413 | -34.3724 | -15.1729
8.7474 | -34.3727 | -15.1823
8.7535 | -34.3732 | -15.1922
8.7596 | -34.3735 | -15.2023
8.7657 | -34.3739 | -15.212
8.7718 | -34.3742 | -15.2215
8.7779 | -34.3746 | -15.2317
8.784 | -34.3748 | -15.2412
8.7901 | -34.3752 | -15.2513
8.7962 | -34.3756 | -15.261
8.8023 | -34.376 | -15.2704
8.8084 | -34.3763 | -15.2803
8.8146 | -34.3766 | -15.2898
8.8207 | -34.377 | -15.2995
8.8268 | -34.3774 | -15.3091
8.8329 | -34.3777 | -15.3189
8.839 | -34.3781 | -15.3287
8.8451 | -34.3785 | -15.3382
8.8512 | -34.3789 | -15.3479
8.8573 | -34.3792 | -15.3575
8.8634 | -34.3796 | -15.367
8.8695 | -34.3799 | -15.3769
8.8756 | -34.3803 | -15.3864
8.8817 | -34.3807 | -15.3963
8.8878 | -34.381 | -15.4062
8.8939 | -34.3813 | -15.4157
8.9 | -34.3817 | -15.4252
8.9061 | -34.382 | -15.4349
8.9122 | -34.3824 | -15.4445
8.9183 | -34.3828 | -15.4541
8.9244 | -34.3831 | -15.4637
8.9305 | -34.3835 | -15.4733
8.9366 | -34.3839 | -15.4829
8.9427 | -34.3843 | -15.4926
8.9488 | -34.3846 | -15.5022
8.955 | -34.385 | -15.5118
8.9611 | -34.3854 | -15.5216
8.9672 | -34.3858 | -15.5316
8.9733 | -34.3861 | -15.5411
8.9794 | -34.3865 | -15.5509
8.9855 | -34.3869 | -15.5605
8.9916 | -34.3872 | -15.5702
8.9977 | -34.3876 | -15.5798
9.0038 | -34.388 | -15.5895
9.0099 | -34.3884 | -15.5993
9.016 | -34.3888 | -15.609
9.0221 | -34.3892 | -15.6188
9.0282 | -34.3895 | -15.6286
9.0343 | -34.3899 | -15.6385
9.0404 | -34.3903 | -15.648
9.0465 | -34.3907 | -15.6581
9.0526 | -34.3911 | -15.6677
9.0587 | -34.3915 | -15.6773
9.0648 | -34.3919 | -15.687
9.0709 | -34.3923 | -15.6964
9.077 | -34.3926 | -15.7067
9.0831 | -34.393 | -15.7164
9.0892 | -34.3934 | -15.7261
9.0953 | -34.3937 | -15.7359
9.1015 | -34.3941 | -15.7456
9.1076 | -34.3945 | -15.7554
9.1137 | -34.3949 | -15.765
9.1198 | -34.3953 | -15.7744
9.1259 | -34.3957 | -15.784
9.132 | -34.3961 | -15.794
9.1381 | -34.3965 | -15.8036
9.1442 | -34.397 | -15.8128
9.1503 | -34.3973 | -15.8226
9.1564 | -34.3977 | -15.8323
9.1625 | -34.3981 | -15.842
9.1686 | -34.3985 | -15.8517
9.1747 | -34.3989 | -15.8615
9.1808 | -34.3993 | -15.8715
9.1869 | -34.3996 | -15.8813
9.193 | -34.4 | -15.8911
9.1991 | -34.4004 | -15.9005
9.2052 | -34.4008 | -15.9101
9.2113 | -34.4012 | -15.92
9.2174 | -34.4016 | -15.9302
9.2235 | -34.402 | -15.9396
9.2296 | -34.4025 | -15.9494
9.2357 | -34.4028 | -15.9592
9.2419 | -34.4032 | -15.9689
9.248 | -34.4037 | -15.9788
9.2541 | -34.4041 | -15.9886
9.2602 | -34.4044 | -15.9984
9.2663 | -34.4048 | -16.0079
9.2724 | -34.4052 | -16.0176
9.2785 | -34.4056 | -16.0273
9.2846 | -34.406 | -16.037
9.2907 | -34.4064 | -16.0466
9.2968 | -34.4068 | -16.0565
9.3029 | -34.4071 | -16.0662
9.309 | -34.4075 | -16.0759
9.3151 | -34.4079 | -16.0855
9.3212 | -34.4084 | -16.0952
9.3273 | -34.4088 | -16.1048
9.3334 | -34.4093 | -16.1143
9.3395 | -34.4097 | -16.1239
9.3456 | -34.4101 | -16.1334
9.3517 | -34.4105 | -16.1427
9.3578 | -34.4109 | -16.1522
9.3639 | -34.4113 | -16.162
9.37 | -34.4117 | -16.1717
9.3761 | -34.4121 | -16.1813
9.3822 | -34.4125 | -16.1909
9.3884 | -34.413 | -16.2005
9.3945 | -34.4134 | -16.2103
9.4006 | -34.4138 | -16.22
9.4067 | -34.4142 | -16.2298
9.4128 | -34.4146 | -16.2394
9.4189 | -34.415 | -16.2492
9.425 | -34.4154 | -16.2587
9.4311 | -34.4158 | -16.2684
9.4372 | -34.4162 | -16.278
9.4433 | -34.4166 | -16.2878
9.4494 | -34.417 | -16.2973
9.4555 | -34.4175 | -16.307
9.4616 | -34.4179 | -16.3166
9.4677 | -34.4184 | -16.3261
9.4738 | -34.4187 | -16.3361
9.4799 | -34.4191 | -16.3457
9.486 | -34.4195 | -16.3556
9.4921 | -34.4199 | -16.3655
9.4982 | -34.4203 | -16.3748
9.5043 | -34.4207 | -16.3842
9.5104 | -34.4212 | -16.3937
9.5165 | -34.4216 | -16.4034
9.5226 | -34.422 | -16.4131
9.5288 | -34.4224 | -16.4227
9.5349 | -34.4229 | -16.4323
9.541 | -34.4233 | -16.4419
9.5471 | -34.4237 | -16.4517
9.5532 | -34.4241 | -16.4613
9.5593 | -34.4246 | -16.471
9.5654 | -34.425 | -16.4808
9.5715 | -34.4254 | -16.4906
9.5776 | -34.4258 | -16.5005
9.5837 | -34.4262 | -16.5099
9.5898 | -34.4265 | -16.5198
9.5959 | -34.4269 | -16.5293
9.602 | -34.4273 | -16.539
9.6081 | -34.4277 | -16.5488
9.6142 | -34.4282 | -16.5583
9.6203 | -34.4286 | -16.568
9.6264 | -34.4289 | -16.5773
9.6325 | -34.4293 | -16.5871
9.6386 | -34.4297 | -16.5967
9.6447 | -34.4301 | -16.6063
9.6508 | -34.4306 | -16.616
9.6569 | -34.431 | -16.6257
9.663 | -34.4314 | -16.6354
9.6691 | -34.4319 | -16.645
9.6753 | -34.4323 | -16.6546
9.6814 | -34.4327 | -16.664
9.6875 | -34.4332 | -16.6734
9.6936 | -34.4336 | -16.6829
9.6997 | -34.434 | -16.6928
9.7058 | -34.4345 | -16.7025
9.7119 | -34.4349 | -16.7121
9.718 | -34.4353 | -16.722
9.7241 | -34.4357 | -16.7316
9.7302 | -34.4361 | -16.7412
9.7363 | -34.4365 | -16.7506
9.7424 | -34.437 | -16.7602
9.7485 | -34.4373 | -16.7697
9.7546 | -34.4378 | -16.7791
9.7607 | -34.4382 | -16.7889
9.7668 | -34.4386 | -16.7988
9.7729 | -34.4391 | -16.8082
9.779 | -34.4395 | -16.8178
9.7851 | -34.4399 | -16.8276
9.7912 | -34.4403 | -16.8376
9.7973 | -34.4407 | -16.8474
9.8034 | -34.4412 | -16.8571
9.8095 | -34.4416 | -16.8667
9.8157 | -34.442 | -16.8764
9.8218 | -34.4424 | -16.8857
9.8279 | -34.4429 | -16.8954
9.834 | -34.4433 | -16.9049
9.8401 | -34.4437 | -16.9145
9.8462 | -34.4441 | -16.9238
9.8523 | -34.4446 | -16.9333
9.8584 | -34.445 | -16.9431
9.8645 | -34.4454 | -16.9527
9.8706 | -34.4458 | -16.9623
9.8767 | -34.4462 | -16.9715
9.8828 | -34.4467 | -16.9807
9.8889 | -34.4471 | -16.9904
9.895 | -34.4475 | -16.9997
9.9011 | -34.4479 | -17.009
9.9072 | -34.4483 | -17.0183
9.9133 | -34.4487 | -17.0277
9.9194 | -34.4492 | -17.0371
9.9255 | -34.4496 | -17.0467
9.9316 | -34.45 | -17.0561
9.9377 | -34.4505 | -17.0654
9.9438 | -34.4509 | -17.0749
9.9499 | -34.4513 | -17.0846
9.956 | -34.4518 | -17.094
9.9622 | -34.4522 | -17.1035
9.9683 | -34.4526 | -17.113
9.9744 | -34.453 | -17.1225
9.9805 | -34.4535 | -17.1321
9.9866 | -34.4539 | -17.1414
9.9927 | -34.4543 | -17.1511
9.9988 | -34.4548 | -17.1606
10.0049 | -34.4552 | -17.1701
10.011 | -34.4556 | -17.1796
10.0171 | -34.456 | -17.1891
10.0232 | -34.4565 | -17.1984
10.0293 | -34.457 | -17.2076
10.0354 | -34.4574 | -17.217
10.0415 | -34.4579 | -17.2264
10.0476 | -34.4583 | -17.2362
10.0537 | -34.4587 | -17.2459
10.0598 | -34.4591 | -17.2556
10.0659 | -34.4595 | -17.265
10.072 | -34.46 | -17.2744
10.0781 | -34.4604 | -17.2838
10.0842 | -34.4608 | -17.2934
10.0903 | -34.4613 | -17.3031
10.0964 | -34.4617 | -17.3124
10.1026 | -34.4622 | -17.3219
10.1087 | -34.4626 | -17.3313
10.1148 | -34.463 | -17.3409
10.1209 | -34.4635 | -17.3502
10.127 | -34.4639 | -17.3594
10.1331 | -34.4644 | -17.3687
10.1392 | -34.4648 | -17.3781
10.1453 | -34.4652 | -17.3875
10.1514 | -34.4656 | -17.3967
10.1575 | -34.466 | -17.4061
10.1636 | -34.4665 | -17.4149
10.1697 | -34.4669 | -17.4244
10.1758 | -34.4674 | -17.4337
10.1819 | -34.4678 | -17.4431
10.188 | -34.4682 | -17.4523
10.1941 | -34.4686 | -17.4618
10.2002 | -34.469 | -17.4708
10.2063 | -34.4695 | -17.4804
10.2124 | -34.4699 | -17.4897
10.2185 | -34.4703 | -17.4991
10.2246 | -34.4708 | -17.5085
10.2307 | -34.4712 | -17.518
10.2368 | -34.4716 | -17.5268
10.2429 | -34.4721 | -17.536
10.2491 | -34.4725 | -17.5451
10.2552 | -34.4729 | -17.5549
10.2613 | -34.4734 | -17.5645
10.2674 | -34.4738 | -17.5741
10.2735 | -34.4742 | -17.5837
10.2796 | -34.4746 | -17.5931
10.2857 | -34.475 | -17.6026
10.2918 | -34.4755 | -17.612
10.2979 | -34.4759 | -17.6216
10.304 | -34.4764 | -17.6311
10.3101 | -34.4768 | -17.6406
10.3162 | -34.4773 | -17.65
10.3223 | -34.4777 | -17.6594
10.3284 | -34.4781 | -17.6687
10.3345 | -34.4785 | -17.678
10.3406 | -34.479 | -17.6872
10.3467 | -34.4794 | -17.6965
10.3528 | -34.4798 | -17.7056
10.3589 | -34.4802 | -17.7151
10.365 | -34.4806 | -17.7246
10.3711 | -34.481 | -17.7342
10.3772 | -34.4814 | -17.7436
10.3833 | -34.4819 | -17.7529
10.3895 | -34.4823 | -17.7622
10.3956 | -34.4827 | -17.7712
10.4017 | -34.4832 | -17.7805
10.4078 | -34.4836 | -17.7899
10.4139 | -34.484 | -17.799
10.42 | -34.4845 | -17.8085
10.4261 | -34.485 | -17.8177
10.4322 | -34.4854 | -17.827
10.4383 | -34.4858 | -17.8364
10.4444 | -34.4862 | -17.8454
10.4505 | -34.4867 | -17.8548
10.4566 | -34.4871 | -17.864
10.4627 | -34.4875 | -17.8732
10.4688 | -34.488 | -17.8822
10.4749 | -34.4884 | -17.8914
10.481 | -34.4888 | -17.9008
10.4871 | -34.4893 | -17.9102
10.4932 | -34.4898 | -17.9194
10.4993 | -34.4902 | -17.9286
10.5054 | -34.4906 | -17.9379
10.5115 | -34.4911 | -17.9472
10.5176 | -34.4916 | -17.9563
10.5237 | -34.492 | -17.9656
10.5298 | -34.4924 | -17.9745
10.536 | -34.4929 | -17.984
10.5421 | -34.4933 | -17.9928
10.5482 | -34.4937 | -18.0023
10.5543 | -34.4942 | -18.0114
10.5604 | -34.4946 | -18.0203
10.5665 | -34.4951 | -18.0294
10.5726 | -34.4955 | -18.0385
10.5787 | -34.4959 | -18.0478
10.5848 | -34.4964 | -18.0566
10.5909 | -34.4968 | -18.0659
10.597 | -34.4972 | -18.0751
10.6031 | -34.4977 | -18.0844
10.6092 | -34.4981 | -18.0939
10.6153 | -34.4985 | -18.1033
10.6214 | -34.499 | -18.1123
10.6275 | -34.4994 | -18.1214
10.6336 | -34.4999 | -18.131
10.6397 | -34.5003 | -18.1401
10.6458 | -34.5007 | -18.1491
10.6519 | -34.5011 | -18.1584
10.658 | -34.5015 | -18.1674
10.6641 | -34.5019 | -18.1768
10.6702 | -34.5024 | -18.1861
10.6764 | -34.5029 | -18.1954
10.6825 | -34.5033 | -18.2046
10.6886 | -34.5038 | -18.2138
10.6947 | -34.5042 | -18.2228
10.7008 | -34.5047 | -18.232
10.7069 | -34.5051 | -18.2412
10.713 | -34.5056 | -18.2503
10.7191 | -34.506 | -18.2597
10.7252 | -34.5065 | -18.2686
10.7313 | -34.507 | -18.2777
10.7374 | -34.5074 | -18.2867
10.7435 | -34.5078 | -18.2961
10.7496 | -34.5082 | -18.3052
10.7557 | -34.5087 | -18.3145
10.7618 | -34.5091 | -18.3237
10.7679 | -34.5096 | -18.3328
10.774 | -34.51 | -18.342
10.7801 | -34.5104 | -18.3515
10.7862 | -34.5109 | -18.3605
10.7923 | -34.5113 | -18.3697
10.7984 | -34.5118 | -18.3787
10.8045 | -34.5122 | -18.388
10.8106 | -34.5126 | -18.3971
10.8168 | -34.5131 | -18.4063
10.8229 | -34.5135 | -18.4155
10.829 | -34.514 | -18.4246
10.8351 | -34.5144 | -18.4337
10.8412 | -34.5148 | -18.4428
10.8473 | -34.5152 | -18.4517
10.8534 | -34.5157 | -18.4612
10.8595 | -34.5161 | -18.4702
10.8656 | -34.5167 | -18.4791
10.8717 | -34.5171 | -18.4879
10.8778 | -34.5176 | -18.4969
10.8839 | -34.518 | -18.5059
10.89 | -34.5184 | -18.5149
10.8961 | -34.5189 | -18.5242
10.9022 | -34.5193 | -18.5332
10.9083 | -34.5198 | -18.5425
10.9144 | -34.5202 | -18.5515
10.9205 | -34.5207 | -18.5606
10.9266 | -34.5212 | -18.57
10.9327 | -34.5217 | -18.5791
10.9388 | -34.5221 | -18.5883
10.9449 | -34.5225 | -18.5975
10.951 | -34.5229 | -18.6065
10.9571 | -34.5234 | -18.6156
10.9633 | -34.5239 | -18.625
10.9694 | -34.5243 | -18.6344
10.9755 | -34.5248 | -18.6435
10.9816 | -34.5252 | -18.6528
10.9877 | -34.5257 | -18.6619
10.9938 | -34.5261 | -18.6712
10.9999 | -34.5266 | -18.6803
11.006 | -34.5271 | -18.6892
11.0121 | -34.5275 | -18.6986
11.0182 | -34.5279 | -18.7075
11.0243 | -34.5284 | -18.7167
11.0304 | -34.5288 | -18.7258
11.0365 | -34.5293 | -18.7349
11.0426 | -34.5297 | -18.7438
11.0487 | -34.5302 | -18.7531
11.0548 | -34.5307 | -18.762
11.0609 | -34.5311 | -18.7714
11.067 | -34.5316 | -18.7805
11.0731 | -34.532 | -18.7897
11.0792 | -34.5324 | -18.7988
11.0853 | -34.5329 | -18.8075
11.0914 | -34.5333 | -18.8165
11.0975 | -34.5338 | -18.8258
11.1037 | -34.5343 | -18.8347
11.1098 | -34.5347 | -18.8441
11.1159 | -34.5352 | -18.8532
11.122 | -34.5356 | -18.8624
11.1281 | -34.5361 | -18.8713
11.1342 | -34.5365 | -18.8802
11.1403 | -34.537 | -18.8896
11.1464 | -34.5374 | -18.8986
11.1525 | -34.5379 | -18.9078
11.1586 | -34.5383 | -18.9173
11.1647 | -34.5388 | -18.9263
11.1708 | -34.5392 | -18.9351
11.1769 | -34.5396 | -18.9438
11.183 | -34.5401 | -18.9528
11.1891 | -34.5405 | -18.9621
11.1952 | -34.541 | -18.9713
11.2013 | -34.5414 | -18.9804
11.2074 | -34.5419 | -18.9892
11.2135 | -34.5423 | -18.9982
11.2196 | -34.5427 | -19.0073
11.2257 | -34.5431 | -19.0163
11.2318 | -34.5435 | -19.0256
11.2379 | -34.5439 | -19.0346
11.244 | -34.5443 | -19.0436
11.2502 | -34.5448 | -19.0526
11.2563 | -34.5453 | -19.0617
11.2624 | -34.5457 | -19.0708
11.2685 | -34.5461 | -19.0797
11.2746 | -34.5466 | -19.0885
11.2807 | -34.547 | -19.0979
11.2868 | -34.5474 | -19.1071
11.2929 | -34.5479 | -19.1158
11.299 | -34.5484 | -19.1252
11.3051 | -34.5488 | -19.1344
11.3112 | -34.5492 | -19.1433
11.3173 | -34.5497 | -19.1524
11.3234 | -34.5501 | -19.1615
11.3295 | -34.5506 | -19.1705
11.3356 | -34.5511 | -19.1795
11.3417 | -34.5515 | -19.1887
11.3478 | -34.5519 | -19.1977
11.3539 | -34.5523 | -19.2067
11.36 | -34.5528 | -19.2156
11.3661 | -34.5532 | -19.2246
11.3722 | -34.5537 | -19.2335
11.3783 | -34.5541 | -19.2426
11.3844 | -34.5545 | -19.2516
11.3906 | -34.5549 | -19.2609
11.3967 | -34.5554 | -19.2697
11.4028 | -34.5558 | -19.279
11.4089 | -34.5562 | -19.2878
11.415 | -34.5567 | -19.2968
11.4211 | -34.5571 | -19.3058
11.4272 | -34.5575 | -19.315
11.4333 | -34.5579 | -19.3241
11.4394 | -34.5584 | -19.3333
11.4455 | -34.5588 | -19.3424
11.4516 | -34.5593 | -19.3513
11.4577 | -34.5597 | -19.3605
11.4638 | -34.5602 | -19.3692
11.4699 | -34.5606 | -19.3787
11.476 | -34.561 | -19.388
11.4821 | -34.5615 | -19.3968
11.4882 | -34.562 | -19.4061
11.4943 | -34.5624 | -19.415
11.5004 | -34.5628 | -19.4242
11.5065 | -34.5633 | -19.4332
11.5126 | -34.5638 | -19.4419
11.5187 | -34.5642 | -19.4509
11.5248 | -34.5647 | -19.46
11.5309 | -34.5651 | -19.4689
11.5371 | -34.5655 | -19.4778
11.5432 | -34.5659 | -19.4869
11.5493 | -34.5664 | -19.496
11.5554 | -34.5668 | -19.5048
11.5615 | -34.5673 | -19.5139
11.5676 | -34.5677 | -19.5231
11.5737 | -34.5682 | -19.532
11.5798 | -34.5687 | -19.5408
11.5859 | -34.5691 | -19.5497
11.592 | -34.5696 | -19.5587
11.5981 | -34.57 | -19.5675
11.6042 | -34.5705 | -19.5766
11.6103 | -34.5709 | -19.5856
11.6164 | -34.5713 | -19.5944
11.6225 | -34.5718 | -19.6033
11.6286 | -34.5722 | -19.6124
11.6347 | -34.5726 | -19.6215
11.6408 | -34.5731 | -19.6304
11.6469 | -34.5736 | -19.6396
11.653 | -34.574 | -19.6484
11.6591 | -34.5745 | -19.6574
11.6652 | -34.575 | -19.6665
11.6713 | -34.5754 | -19.6756
11.6775 | -34.5759 | -19.6844
11.6836 | -34.5763 | -19.6935
11.6897 | -34.5768 | -19.7024
11.6958 | -34.5772 | -19.7118
11.7019 | -34.5777 | -19.7208
11.708 | -34.5781 | -19.73
11.7141 | -34.5786 | -19.7391
11.7202 | -34.5791 | -19.7479
11.7263 | -34.5795 | -19.757
11.7324 | -34.5799 | -19.766
11.7385 | -34.5804 | -19.7752
11.7446 | -34.5808 | -19.7843
11.7507 | -34.5813 | -19.7932
11.7568 | -34.5817 | -19.8023
11.7629 | -34.5822 | -19.8114
11.769 | -34.5827 | -19.8203
11.7751 | -34.5831 | -19.829
11.7812 | -34.5836 | -19.8378
11.7873 | -34.584 | -19.847
11.7934 | -34.5845 | -19.8561
11.7995 | -34.5849 | -19.8648
11.8056 | -34.5854 | -19.8741
11.8117 | -34.5857 | -19.8832
11.8178 | -34.5862 | -19.8921
11.824 | -34.5867 | -19.901
11.8301 | -34.5872 | -19.9096
11.8362 | -34.5876 | -19.9183
11.8423 | -34.588 | -19.9273
11.8484 | -34.5885 | -19.9363
11.8545 | -34.589 | -19.9455
11.8606 | -34.5894 | -19.9545
11.8667 | -34.5899 | -19.9635
11.8728 | -34.5903 | -19.9724
11.8789 | -34.5907 | -19.9818
11.885 | -34.5911 | -19.9906
11.8911 | -34.5916 | -19.9998
11.8972 | -34.5921 | -20.0085
11.9033 | -34.5925 | -20.0171
11.9094 | -34.5929 | -20.0257
11.9155 | -34.5933 | -20.0347
11.9216 | -34.5938 | -20.0436
11.9277 | -34.5943 | -20.0524
11.9338 | -34.5947 | -20.0614
11.9399 | -34.5952 | -20.0706
11.946 | -34.5957 | -20.0796
11.9521 | -34.5961 | -20.0885
11.9582 | -34.5966 | -20.0973
11.9644 | -34.597 | -20.1061
11.9705 | -34.5975 | -20.1153
11.9766 | -34.5979 | -20.1242
11.9827 | -34.5984 | -20.133
11.9888 | -34.5988 | -20.1418
11.9949 | -34.5993 | -20.1507
12.001 | -34.5998 | -20.1596
12.0071 | -34.6002 | -20.1684
12.0132 | -34.6007 | -20.1771
12.0193 | -34.6011 | -20.1862
12.0254 | -34.6016 | -20.195
12.0315 | -34.602 | -20.2038
12.0376 | -34.6025 | -20.2125
12.0437 | -34.6029 | -20.2213
12.0498 | -34.6033 | -20.2304
12.0559 | -34.6038 | -20.2394
12.062 | -34.6043 | -20.2483
12.0681 | -34.6047 | -20.2572
12.0742 | -34.6052 | -20.2663
12.0803 | -34.6056 | -20.2751
12.0864 | -34.6061 | -20.284
12.0925 | -34.6065 | -20.2929
12.0986 | -34.607 | -20.3016
12.1047 | -34.6075 | -20.3106
12.1109 | -34.608 | -20.3195
12.117 | -34.6084 | -20.3283
12.1231 | -34.6089 | -20.3372
12.1292 | -34.6094 | -20.3462
12.1353 | -34.6098 | -20.3549
12.1414 | -34.6103 | -20.3639
12.1475 | -34.6107 | -20.3728
12.1536 | -34.6112 | -20.3817
12.1597 | -34.6116 | -20.3907
12.1658 | -34.6121 | -20.3996
12.1719 | -34.6126 | -20.4085
12.178 | -34.613 | -20.4176
12.1841 | -34.6135 | -20.4265
12.1902 | -34.614 | -20.4354
12.1963 | -34.6144 | -20.4443
12.2024 | -34.6149 | -20.453
12.2085 | -34.6153 | -20.4619
12.2146 | -34.6158 | -20.4708
12.2207 | -34.6163 | -20.4797
12.2268 | -34.6167 | -20.4886
12.2329 | -34.6172 | -20.4974
12.239 | -34.6176 | -20.5059
12.2451 | -34.6181 | -20.5146
12.2513 | -34.6186 | -20.5236
12.2574 | -34.6191 | -20.5327
12.2635 | -34.6195 | -20.5413
12.2696 | -34.62 | -20.55
12.2757 | -34.6204 | -20.5587
12.2818 | -34.6209 | -20.5674
12.2879 | -34.6213 | -20.5761
12.294 | -34.6218 | -20.5848
12.3001 | -34.6222 | -20.5938
12.3062 | -34.6227 | -20.6026
12.3123 | -34.6232 | -20.6115
12.3184 | -34.6237 | -20.6202
12.3245 | -34.6242 | -20.6292
12.3306 | -34.6246 | -20.6378
12.3367 | -34.6251 | -20.6465
12.3428 | -34.6255 | -20.6551
12.3489 | -34.626 | -20.6635
12.355 | -34.6265 | -20.6723
12.3611 | -34.6269 | -20.6811
12.3672 | -34.6274 | -20.6897
12.3733 | -34.6278 | -20.6987
12.3794 | -34.6283 | -20.707
12.3855 | -34.6287 | -20.7157
12.3916 | -34.6292 | -20.7247
12.3978 | -34.6297 | -20.7337
12.4039 | -34.6301 | -20.7425
12.41 | -34.6306 | -20.7514
12.4161 | -34.631 | -20.7599
12.4222 | -34.6315 | -20.7686
12.4283 | -34.632 | -20.7774
12.4344 | -34.6325 | -20.7863
12.4405 | -34.6329 | -20.7951
12.4466 | -34.6333 | -20.8038
12.4527 | -34.6338 | -20.8125
12.4588 | -34.6343 | -20.8213
12.4649 | -34.6347 | -20.8303
12.471 | -34.6352 | -20.8395
12.4771 | -34.6357 | -20.8478
12.4832 | -34.6361 | -20.8566
12.4893 | -34.6366 | -20.8652
12.4954 | -34.6371 | -20.8738
12.5015 | -34.6376 | -20.8823
12.5076 | -34.638 | -20.8912
12.5137 | -34.6385 | -20.9003
12.5198 | -34.6389 | -20.9088
12.5259 | -34.6394 | -20.9175
12.532 | -34.6399 | -20.9264
12.5382 | -34.6404 | -20.9351
12.5443 | -34.6408 | -20.9438
12.5504 | -34.6413 | -20.9526
12.5565 | -34.6418 | -20.9612
12.5626 | -34.6422 | -20.9702
12.5687 | -34.6427 | -20.979
12.5748 | -34.6432 | -20.9876
12.5809 | -34.6436 | -20.9961
12.587 | -34.644 | -21.0051
12.5931 | -34.6445 | -21.0138
12.5992 | -34.645 | -21.0229
12.6053 | -34.6454 | -21.0315
12.6114 | -34.6459 | -21.0402
12.6175 | -34.6464 | -21.0491
12.6236 | -34.6468 | -21.0576
12.6297 | -34.6474 | -21.0663
12.6358 | -34.6478 | -21.0748
12.6419 | -34.6483 | -21.0838
12.648 | -34.6488 | -21.0927
12.6541 | -34.6493 | -21.1014
12.6602 | -34.6498 | -21.1097
12.6663 | -34.6502 | -21.1183
12.6724 | -34.6507 | -21.127
12.6785 | -34.6512 | -21.136
12.6847 | -34.6517 | -21.1447
12.6908 | -34.6521 | -21.1534
12.6969 | -34.6527 | -21.1619
12.703 | -34.6531 | -21.1708
12.7091 | -34.6536 | -21.1797
12.7152 | -34.654 | -21.1883
12.7213 | -34.6545 | -21.1972
12.7274 | -34.655 | -21.206
12.7335 | -34.6555 | -21.2141
12.7396 | -34.6559 | -21.2224
12.7457 | -34.6564 | -21.2311
12.7518 | -34.6569 | -21.2398
12.7579 | -34.6573 | -21.2488
12.764 | -34.6577 | -21.2575
12.7701 | -34.6581 | -21.266
12.7762 | -34.6586 | -21.2747
12.7823 | -34.6591 | -21.2833
12.7884 | -34.6595 | -21.292
12.7945 | -34.66 | -21.3003
12.8006 | -34.6605 | -21.3085
12.8067 | -34.661 | -21.3173
12.8128 | -34.6614 | -21.3257
12.8189 | -34.6619 | -21.334
12.8251 | -34.6623 | -21.3424
12.8312 | -34.6627 | -21.3512
12.8373 | -34.6632 | -21.3597
12.8434 | -34.6637 | -21.3683
12.8495 | -34.6642 | -21.3767
12.8556 | -34.6646 | -21.3852
12.8617 | -34.6651 | -21.3937
12.8678 | -34.6655 | -21.4023
12.8739 | -34.666 | -21.4106
12.88 | -34.6665 | -21.4193
12.8861 | -34.6669 | -21.428
12.8922 | -34.6674 | -21.4367
12.8983 | -34.6679 | -21.4449
12.9044 | -34.6684 | -21.4534
12.9105 | -34.6689 | -21.462
12.9166 | -34.6694 | -21.4708
12.9227 | -34.6698 | -21.4798
12.9288 | -34.6703 | -21.4883
12.9349 | -34.6708 | -21.4972
12.941 | -34.6713 | -21.5058
12.9471 | -34.6718 | -21.5144
12.9532 | -34.6723 | -21.523
12.9593 | -34.6728 | -21.5318
12.9654 | -34.6732 | -21.5406
12.9716 | -34.6737 | -21.5493
12.9777 | -34.6741 | -21.5577
12.9838 | -34.6746 | -21.5661
12.9899 | -34.675 | -21.5749
12.996 | -34.6755 | -21.5837
13.0021 | -34.6759 | -21.5924
13.0082 | -34.6764 | -21.601
13.0143 | -34.6768 | -21.6092
13.0204 | -34.6773 | -21.6174
13.0265 | -34.6778 | -21.6258
13.0326 | -34.6782 | -21.6346
13.0387 | -34.6787 | -21.6434
13.0448 | -34.6792 | -21.6519
13.0509 | -34.6797 | -21.6606
13.057 | -34.6802 | -21.6689
13.0631 | -34.6807 | -21.6773
13.0692 | -34.6812 | -21.6858
13.0753 | -34.6817 | -21.6944
13.0814 | -34.6821 | -21.7029
13.0875 | -34.6826 | -21.7112
13.0936 | -34.6831 | -21.7197
13.0997 | -34.6835 | -21.7281
13.1058 | -34.684 | -21.7362
13.112 | -34.6844 | -21.7447
13.1181 | -34.6849 | -21.7533
13.1242 | -34.6853 | -21.7619
13.1303 | -34.6858 | -21.7705
13.1364 | -34.6862 | -21.779
13.1425 | -34.6867 | -21.7875
13.1486 | -34.6872 | -21.7959
13.1547 | -34.6877 | -21.8042
13.1608 | -34.6881 | -21.8126
13.1669 | -34.6886 | -21.821
13.173 | -34.6891 | -21.8294
13.1791 | -34.6895 | -21.8379
13.1852 | -34.69 | -21.8467
13.1913 | -34.6905 | -21.8553
13.1974 | -34.6909 | -21.8637
13.2035 | -34.6913 | -21.8723
13.2096 | -34.6919 | -21.8808
13.2157 | -34.6923 | -21.8893
13.2218 | -34.6928 | -21.8979
13.2279 | -34.6933 | -21.9065
13.234 | -34.6938 | -21.915
13.2401 | -34.6943 | -21.9235
13.2462 | -34.6947 | -21.9318
13.2524 | -34.6952 | -21.9402
13.2585 | -34.6957 | -21.9486
13.2646 | -34.6961 | -21.9574
13.2707 | -34.6966 | -21.966
13.2768 | -34.6971 | -21.9748
13.2829 | -34.6975 | -21.9833
13.289 | -34.698 | -21.9918
13.2951 | -34.6985 | -22.0001
13.3012 | -34.6989 | -22.0087
13.3073 | -34.6994 | -22.0173
13.3134 | -34.6998 | -22.0257
13.3195 | -34.7003 | -22.034
13.3256 | -34.7008 | -22.0427
13.3317 | -34.7013 | -22.0511
13.3378 | -34.7018 | -22.0597
13.3439 | -34.7022 | -22.0682
13.35 | -34.7027 | -22.0767
13.3561 | -34.7032 | -22.0852
13.3622 | -34.7036 | -22.0936
13.3683 | -34.7041 | -22.1016
13.3744 | -34.7045 | -22.1099
13.3805 | -34.705 | -22.1185
13.3866 | -34.7054 | -22.1266
13.3927 | -34.7059 | -22.1351
13.3989 | -34.7063 | -22.1433
13.405 | -34.7068 | -22.1517
13.4111 | -34.7073 | -22.1601
13.4172 | -34.7078 | -22.1683
13.4233 | -34.7082 | -22.177
13.4294 | -34.7087 | -22.1853
13.4355 | -34.7092 | -22.1943
13.4416 | -34.7096 | -22.2028
13.4477 | -34.71 | -22.2113
13.4538 | -34.7105 | -22.2197
13.4599 | -34.711 | -22.2283
13.466 | -34.7114 | -22.237
13.4721 | -34.7119 | -22.2456
13.4782 | -34.7124 | -22.2541
13.4843 | -34.7129 | -22.2625
13.4904 | -34.7134 | -22.2706
13.4965 | -34.7138 | -22.2793
13.5026 | -34.7143 | -22.2877
13.5087 | -34.7148 | -22.2957
13.5148 | -34.7152 | -22.3043
13.5209 | -34.7157 | -22.3126
13.527 | -34.7161 | -22.3213
13.5331 | -34.7165 | -22.3296
13.5393 | -34.717 | -22.3381
13.5454 | -34.7175 | -22.3464
13.5515 | -34.718 | -22.3549
13.5576 | -34.7185 | -22.3635
13.5637 | -34.719 | -22.3718
13.5698 | -34.7194 | -22.3804
13.5759 | -34.7199 | -22.3885
13.582 | -34.7203 | -22.397
13.5881 | -34.7208 | -22.4056
13.5942 | -34.7212 | -22.4141
13.6003 | -34.7218 | -22.4226
13.6064 | -34.7222 | -22.4312
13.6125 | -34.7227 | -22.4399
13.6186 | -34.7232 | -22.4483
13.6247 | -34.7237 | -22.4568
13.6308 | -34.7241 | -22.4649
13.6369 | -34.7246 | -22.4733
13.643 | -34.7251 | -22.4816
13.6491 | -34.7256 | -22.4898
13.6552 | -34.726 | -22.4981
13.6613 | -34.7265 | -22.5068
13.6674 | -34.727 | -22.5154
13.6735 | -34.7274 | -22.5238
13.6796 | -34.7279 | -22.532
13.6858 | -34.7283 | -22.5405
13.6919 | -34.7288 | -22.5489
13.698 | -34.7293 | -22.5574
13.7041 | -34.7297 | -22.566
13.7102 | -34.7302 | -22.5745
13.7163 | -34.7307 | -22.5832
13.7224 | -34.7312 | -22.5915
13.7285 | -34.7317 | -22.5999
13.7346 | -34.7321 | -22.6082
13.7407 | -34.7326 | -22.6166
13.7468 | -34.733 | -22.6251
13.7529 | -34.7335 | -22.6334
13.759 | -34.7339 | -22.6421
13.7651 | -34.7344 | -22.6507
13.7712 | -34.7349 | -22.6594
13.7773 | -34.7353 | -22.6679
13.7834 | -34.7358 | -22.6762
13.7895 | -34.7363 | -22.6845
13.7956 | -34.7367 | -22.6931
13.8017 | -34.7372 | -22.7013
13.8078 | -34.7377 | -22.7097
13.8139 | -34.7382 | -22.718
13.82 | -34.7386 | -22.7261
13.8262 | -34.7391 | -22.7347
13.8323 | -34.7396 | -22.7434
13.8384 | -34.7401 | -22.7521
13.8445 | -34.7406 | -22.7606
13.8506 | -34.741 | -22.769
13.8567 | -34.7415 | -22.7776
13.8628 | -34.742 | -22.7861
13.8689 | -34.7424 | -22.7945
13.875 | -34.7429 | -22.8028
13.8811 | -34.7434 | -22.8111
13.8872 | -34.7439 | -22.8195
13.8933 | -34.7444 | -22.8276
13.8994 | -34.7448 | -22.8357
13.9055 | -34.7453 | -22.8442
13.9116 | -34.7458 | -22.8524
13.9177 | -34.7463 | -22.8606
13.9238 | -34.7468 | -22.8691
13.9299 | -34.7473 | -22.8775
13.936 | -34.7477 | -22.8859
13.9421 | -34.7482 | -22.8943
13.9482 | -34.7486 | -22.9028
13.9543 | -34.7491 | -22.911
13.9604 | -34.7496 | -22.9194
13.9665 | -34.7501 | -22.928
13.9727 | -34.7506 | -22.9364
13.9788 | -34.751 | -22.945
13.9849 | -34.7515 | -22.9534
13.991 | -34.7519 | -22.962
13.9971 | -34.7524 | -22.9707
14.0032 | -34.7528 | -22.9792
14.0093 | -34.7533 | -22.9876
14.0154 | -34.7537 | -22.9961
14.0215 | -34.7542 | -23.0044
14.0276 | -34.7547 | -23.0126
14.0337 | -34.7552 | -23.0212
14.0398 | -34.7557 | -23.0294
14.0459 | -34.7561 | -23.0379
14.052 | -34.7566 | -23.0463
14.0581 | -34.757 | -23.0544
14.0642 | -34.7575 | -23.0629
14.0703 | -34.758 | -23.0713
14.0764 | -34.7584 | -23.0799
14.0825 | -34.7589 | -23.0881
14.0886 | -34.7594 | -23.0962
14.0947 | -34.7598 | -23.1044
14.1008 | -34.7603 | -23.1131
14.1069 | -34.7607 | -23.1212
14.1131 | -34.7612 | -23.1299
14.1192 | -34.7616 | -23.1379
14.1253 | -34.7621 | -23.1462
14.1314 | -34.7626 | -23.1546
14.1375 | -34.7631 | -23.1631
14.1436 | -34.7636 | -23.1714
14.1497 | -34.7641 | -23.1796
14.1558 | -34.7645 | -23.1881
14.1619 | -34.765 | -23.1964
14.168 | -34.7655 | -23.2048
14.1741 | -34.766 | -23.2131
14.1802 | -34.7664 | -23.2212
14.1863 | -34.7669 | -23.2297
14.1924 | -34.7674 | -23.238
14.1985 | -34.7679 | -23.2462
14.2046 | -34.7683 | -23.2545
14.2107 | -34.7688 | -23.2625
14.2168 | -34.7693 | -23.2712
14.2229 | -34.7698 | -23.2795
14.229 | -34.7702 | -23.288
14.2351 | -34.7707 | -23.2961
14.2412 | -34.7712 | -23.3044
14.2473 | -34.7717 | -23.3128
14.2534 | -34.7722 | -23.321
14.2596 | -34.7726 | -23.3294
14.2657 | -34.7731 | -23.3377
14.2718 | -34.7736 | -23.3462
14.2779 | -34.774 | -23.3544
14.284 | -34.7745 | -23.3627
14.2901 | -34.775 | -23.3709
14.2962 | -34.7756 | -23.3791
14.3023 | -34.776 | -23.3875
14.3084 | -34.7765 | -23.3961
14.3145 | -34.777 | -23.4044
14.3206 | -34.7774 | -23.4128
14.3267 | -34.7779 | -23.4211
14.3328 | -34.7784 | -23.4293
14.3389 | -34.7788 | -23.4374
14.345 | -34.7793 | -23.4459
14.3511 | -34.7798 | -23.454
14.3572 | -34.7803 | -23.4625
14.3633 | -34.7807 | -23.4708
14.3694 | -34.7812 | -23.479
14.3755 | -34.7816 | -23.487
14.3816 | -34.7822 | -23.4953
14.3877 | -34.7827 | -23.5034
14.3938 | -34.7831 | -23.5119
14.4 | -34.7836 | -23.52
14.4061 | -34.7841 | -23.5284
14.4122 | -34.7845 | -23.5367
14.4183 | -34.7849 | -23.545
14.4244 | -34.7854 | -23.5532
14.4305 | -34.7859 | -23.5617
14.4366 | -34.7864 | -23.5701
14.4427 | -34.7869 | -23.5787
14.4488 | -34.7873 | -23.5868
14.4549 | -34.7878 | -23.5949
14.461 | -34.7883 | -23.6032
14.4671 | -34.7888 | -23.6115
14.4732 | -34.7893 | -23.6195
14.4793 | -34.7897 | -23.6284
14.4854 | -34.7903 | -23.6364
14.4915 | -34.7907 | -23.6444
14.4976 | -34.7913 | -23.6526
14.5037 | -34.7918 | -23.661
14.5098 | -34.7922 | -23.6691
14.5159 | -34.7927 | -23.6776
14.522 | -34.7931 | -23.6859
14.5281 | -34.7936 | -23.6939
14.5342 | -34.7941 | -23.7023
14.5403 | -34.7946 | -23.7106
14.5465 | -34.7951 | -23.7191
14.5526 | -34.7956 | -23.7276
14.5587 | -34.796 | -23.7358
14.5648 | -34.7965 | -23.7439
14.5709 | -34.797 | -23.7523
14.577 | -34.7974 | -23.7605
14.5831 | -34.7979 | -23.7688
14.5892 | -34.7985 | -23.7767
14.5953 | -34.7989 | -23.7846
14.6014 | -34.7994 | -23.7928
14.6075 | -34.7999 | -23.8008
14.6136 | -34.8004 | -23.809
14.6197 | -34.8008 | -23.8173
14.6258 | -34.8013 | -23.8256
14.6319 | -34.8018 | -23.8336
14.638 | -34.8023 | -23.8418
14.6441 | -34.8028 | -23.8497
14.6502 | -34.8032 | -23.8577
14.6563 | -34.8037 | -23.8658
14.6624 | -34.8042 | -23.8738
14.6685 | -34.8047 | -23.8817
14.6746 | -34.8052 | -23.8899
14.6807 | -34.8057 | -23.8981
14.6869 | -34.8062 | -23.9062
14.693 | -34.8067 | -23.9143
14.6991 | -34.8071 | -23.9223
14.7052 | -34.8076 | -23.9305
14.7113 | -34.8081 | -23.9388
14.7174 | -34.8086 | -23.947
14.7235 | -34.809 | -23.955
14.7296 | -34.8095 | -23.9631
14.7357 | -34.81 | -23.9711
14.7418 | -34.8105 | -23.979
14.7479 | -34.811 | -23.9872
14.754 | -34.8115 | -23.9953
14.7601 | -34.812 | -24.0035
14.7662 | -34.8124 | -24.0116
14.7723 | -34.8129 | -24.0196
14.7784 | -34.8134 | -24.0279
14.7845 | -34.8139 | -24.0361
14.7906 | -34.8144 | -24.0443
14.7967 | -34.8148 | -24.0525
14.8028 | -34.8153 | -24.0605
14.8089 | -34.8158 | -24.0685
14.815 | -34.8163 | -24.0765
14.8211 | -34.8167 | -24.0846
14.8272 | -34.8172 | -24.0928
14.8334 | -34.8177 | -24.1007
14.8395 | -34.8182 | -24.1088
14.8456 | -34.8187 | -24.1168
14.8517 | -34.8191 | -24.1249
14.8578 | -34.8196 | -24.133
14.8639 | -34.8201 | -24.1412
14.87 | -34.8206 | -24.1494
14.8761 | -34.8211 | -24.1572
14.8822 | -34.8216 | -24.1654
14.8883 | -34.8221 | -24.1733
14.8944 | -34.8226 | -24.1815
14.9005 | -34.8231 | -24.1894
14.9066 | -34.8235 | -24.1975
14.9127 | -34.824 | -24.2055
14.9188 | -34.8245 | -24.2137
14.9249 | -34.825 | -24.2218
14.931 | -34.8254 | -24.2298
14.9371 | -34.8259 | -24.2377
14.9432 | -34.8264 | -24.2458
14.9493 | -34.8269 | -24.2538
14.9554 | -34.8274 | -24.2619
14.9615 | -34.8279 | -24.27
14.9676 | -34.8284 | -24.2783
14.9738 | -34.8288 | -24.2863
14.9799 | -34.8293 | -24.2948
14.986 | -34.8298 | -24.3025
14.9921 | -34.8303 | -24.3105
14.9982 | -34.8308 | -24.3187
15.0043 | -34.8312 | -24.3268
15.0104 | -34.8317 | -24.3349
15.0165 | -34.8323 | -24.3431
15.0226 | -34.8328 | -24.3508
15.0287 | -34.8332 | -24.3592
15.0348 | -34.8338 | -24.3674
15.0409 | -34.8342 | -24.3757
15.047 | -34.8347 | -24.384
15.0531 | -34.8353 | -24.392
15.0592 | -34.8358 | -24.4003
15.0653 | -34.8363 | -24.4083
15.0714 | -34.8368 | -24.4162
15.0775 | -34.8373 | -24.424
15.0836 | -34.8378 | -24.4321
15.0897 | -34.8383 | -24.44
15.0958 | -34.8388 | -24.4479
15.1019 | -34.8393 | -24.4559
15.108 | -34.8398 | -24.464
15.1141 | -34.8403 | -24.4721
15.1203 | -34.8408 | -24.4801
15.1264 | -34.8413 | -24.4881
15.1325 | -34.8418 | -24.4961
15.1386 | -34.8423 | -24.504
15.1447 | -34.8429 | -24.5121
15.1508 | -34.8434 | -24.5203
15.1569 | -34.8439 | -24.5285
15.163 | -34.8445 | -24.536
15.1691 | -34.845 | -24.544
15.1752 | -34.8455 | -24.5518
15.1813 | -34.846 | -24.5598
15.1874 | -34.8465 | -24.5679
15.1935 | -34.847 | -24.5755
15.1996 | -34.8475 | -24.5838
15.2057 | -34.848 | -24.5915
15.2118 | -34.8485 | -24.5995
15.2179 | -34.849 | -24.6074
15.224 | -34.8495 | -24.6155
15.2301 | -34.85 | -24.6231
15.2362 | -34.8505 | -24.631
15.2423 | -34.851 | -24.6393
15.2484 | -34.8515 | -24.6471
15.2545 | -34.852 | -24.655
15.2607 | -34.8526 | -24.6631
15.2668 | -34.8531 | -24.6708
15.2729 | -34.8536 | -24.6785
15.279 | -34.8541 | -24.6865
15.2851 | -34.8546 | -24.6944
15.2912 | -34.8551 | -24.7023
15.2973 | -34.8556 | -24.7101
15.3034 | -34.8562 | -24.7178
15.3095 | -34.8566 | -24.7256
15.3156 | -34.8571 | -24.7336
15.3217 | -34.8576 | -24.7413
15.3278 | -34.8582 | -24.7494
15.3339 | -34.8587 | -24.7571
15.34 | -34.8592 | -24.7649
15.3461 | -34.8597 | -24.7729
15.3522 | -34.8602 | -24.7807
15.3583 | -34.8607 | -24.7886
15.3644 | -34.8612 | -24.7965
15.3705 | -34.8618 | -24.8042
15.3766 | -34.8623 | -24.8124
15.3827 | -34.8628 | -24.8202
15.3888 | -34.8633 | -24.8279
15.3949 | -34.8638 | -24.8358
15.401 | -34.8643 | -24.8438
15.4072 | -34.8648 | -24.8514
15.4133 | -34.8653 | -24.8592
15.4194 | -34.8659 | -24.8672
15.4255 | -34.8663 | -24.875
15.4316 | -34.8669 | -24.8829
15.4377 | -34.8673 | -24.8908
15.4438 | -34.8679 | -24.8983
15.4499 | -34.8684 | -24.906
15.456 | -34.8688 | -24.9139
15.4621 | -34.8694 | -24.9218
15.4682 | -34.8699 | -24.9298
15.4743 | -34.8705 | -24.9375
15.4804 | -34.871 | -24.9455
15.4865 | -34.8714 | -24.9532
15.4926 | -34.872 | -24.9611
15.4987 | -34.8725 | -24.969
15.5048 | -34.873 | -24.9769
15.5109 | -34.8735 | -24.9843
15.517 | -34.8741 | -24.9923
15.5231 | -34.8746 | -25.0002
15.5292 | -34.8751 | -25.0078
15.5353 | -34.8756 | -25.0156
15.5414 | -34.8761 | -25.0236
15.5476 | -34.8767 | -25.0315
15.5537 | -34.8772 | -25.0393
15.5598 | -34.8777 | -25.0469
15.5659 | -34.8782 | -25.0549
15.572 | -34.8788 | -25.0626
15.5781 | -34.8792 | -25.0705
15.5842 | -34.8798 | -25.0779
15.5903 | -34.8803 | -25.0856
15.5964 | -34.8808 | -25.0932
15.6025 | -34.8813 | -25.101
15.6086 | -34.8818 | -25.1088
15.6147 | -34.8823 | -25.1166
15.6208 | -34.8828 | -25.1242
15.6269 | -34.8834 | -25.1318
15.633 | -34.8839 | -25.1396
15.6391 | -34.8844 | -25.1475
15.6452 | -34.8849 | -25.1553
15.6513 | -34.8854 | -25.1631
15.6574 | -34.886 | -25.1707
15.6635 | -34.8865 | -25.1785
15.6696 | -34.887 | -25.1862
15.6757 | -34.8875 | -25.1943
15.6818 | -34.888 | -25.2022
15.688 | -34.8886 | -25.2098
15.6941 | -34.8891 | -25.2176
15.7002 | -34.8896 | -25.2252
15.7063 | -34.8901 | -25.2331
15.7124 | -34.8906 | -25.241
15.7185 | -34.8911 | -25.2488
15.7246 | -34.8916 | -25.2564
15.7307 | -34.8922 | -25.2644
15.7368 | -34.8927 | -25.2723
15.7429 | -34.8932 | -25.28
15.749 | -34.8938 | -25.2877
15.7551 | -34.8942 | -25.2955
15.7612 | -34.8947 | -25.3027
15.7673 | -34.8952 | -25.3103
15.7734 | -34.8957 | -25.3179
15.7795 | -34.8962 | -25.3259
15.7856 | -34.8968 | -25.3334
15.7917 | -34.8973 | -25.3408
15.7978 | -34.8978 | -25.3485
15.8039 | -34.8983 | -25.3563
15.81 | -34.8988 | -25.3639
15.8161 | -34.8993 | -25.3718
15.8222 | -34.8998 | -25.3795
15.8283 | -34.9004 | -25.3871
15.8345 | -34.9009 | -25.3949
15.8406 | -34.9015 | -25.4026
15.8467 | -34.902 | -25.4104
15.8528 | -34.9025 | -25.4182
15.8589 | -34.903 | -25.4258
15.865 | -34.9035 | -25.4333
15.8711 | -34.9041 | -25.4412
15.8772 | -34.9047 | -25.4489
15.8833 | -34.9052 | -25.457
15.8894 | -34.9057 | -25.4647
15.8955 | -34.9062 | -25.4724
15.9016 | -34.9068 | -25.4802
15.9077 | -34.9073 | -25.4879
15.9138 | -34.9078 | -25.4955
15.9199 | -34.9083 | -25.5033
15.926 | -34.9088 | -25.5113
15.9321 | -34.9094 | -25.519
15.9382 | -34.9099 | -25.5267
15.9443 | -34.9105 | -25.5345
15.9504 | -34.911 | -25.5422
15.9565 | -34.9116 | -25.55
15.9626 | -34.9121 | -25.5579
15.9687 | -34.9127 | -25.5655
15.9749 | -34.9132 | -25.5733
15.981 | -34.9138 | -25.5809
15.9871 | -34.9142 | -25.5887
15.9932 | -34.9148 | -25.5965
15.9993 | -34.9153 | -25.604
16.0054 | -34.9159 | -25.6113
16.0115 | -34.9164 | -25.6188
16.0176 | -34.9169 | -25.6265
16.0237 | -34.9175 | -25.6345
16.0298 | -34.918 | -25.6417
16.0359 | -34.9185 | -25.6494
16.042 | -34.9191 | -25.657
16.0481 | -34.9196 | -25.6648
16.0542 | -34.9201 | -25.6728
16.0603 | -34.9206 | -25.6803
16.0664 | -34.9212 | -25.6879
16.0725 | -34.9217 | -25.6953
16.0786 | -34.9222 | -25.7032
16.0847 | -34.9227 | -25.7109
16.0908 | -34.9233 | -25.7185
16.0969 | -34.9238 | -25.7259
16.103 | -34.9243 | -25.7337
16.1091 | -34.9248 | -25.7412
16.1152 | -34.9253 | -25.7489
16.1214 | -34.9259 | -25.7565
16.1275 | -34.9264 | -25.764
16.1336 | -34.927 | -25.7715
16.1397 | -34.9275 | -25.7792
16.1458 | -34.9281 | -25.7867
16.1519 | -34.9286 | -25.7946
16.158 | -34.9291 | -25.8023
16.1641 | -34.9296 | -25.81
16.1702 | -34.9301 | -25.8177
16.1763 | -34.9306 | -25.8254
16.1824 | -34.9312 | -25.8328
16.1885 | -34.9318 | -25.8405
16.1946 | -34.9323 | -25.8481
16.2007 | -34.9328 | -25.856
16.2068 | -34.9333 | -25.8634
16.2129 | -34.9339 | -25.8711
16.219 | -34.9344 | -25.8787
16.2251 | -34.9349 | -25.8863
16.2312 | -34.9354 | -25.894
16.2373 | -34.9359 | -25.9017
16.2434 | -34.9365 | -25.9091
16.2495 | -34.937 | -25.9166
16.2556 | -34.9376 | -25.9243
16.2618 | -34.9381 | -25.9317
16.2679 | -34.9387 | -25.9395
16.274 | -34.9392 | -25.9473
16.2801 | -34.9397 | -25.9548
16.2862 | -34.9402 | -25.9624
16.2923 | -34.9408 | -25.9703
16.2984 | -34.9412 | -25.978
16.3045 | -34.9418 | -25.9854
16.3106 | -34.9424 | -25.9928
16.3167 | -34.9429 | -26.0005
16.3228 | -34.9434 | -26.0081
16.3289 | -34.9439 | -26.0158
16.335 | -34.9444 | -26.0235
16.3411 | -34.945 | -26.0309
16.3472 | -34.9455 | -26.0387
16.3533 | -34.946 | -26.0463
16.3594 | -34.9465 | -26.0537
16.3655 | -34.947 | -26.0615
16.3716 | -34.9475 | -26.0689
16.3777 | -34.9481 | -26.0765
16.3838 | -34.9486 | -26.0841
16.3899 | -34.9491 | -26.0917
16.396 | -34.9496 | -26.0993
16.4021 | -34.9502 | -26.1068
16.4083 | -34.9507 | -26.1144
16.4144 | -34.9513 | -26.122
16.4205 | -34.9518 | -26.1297
16.4266 | -34.9523 | -26.1373
16.4327 | -34.9529 | -26.1448
16.4388 | -34.9534 | -26.1526
16.4449 | -34.9539 | -26.16
16.451 | -34.9544 | -26.1678
16.4571 | -34.9549 | -26.1755
16.4632 | -34.9554 | -26.1831
16.4693 | -34.9559 | -26.1908
16.4754 | -34.9565 | -26.1984
16.4815 | -34.957 | -26.2064
16.4876 | -34.9576 | -26.2139
16.4937 | -34.9581 | -26.2214
16.4998 | -34.9587 | -26.229
16.5059 | -34.9592 | -26.2365
16.512 | -34.9598 | -26.2444
16.5181 | -34.9603 | -26.2521
16.5242 | -34.9608 | -26.2597
16.5303 | -34.9613 | -26.2672
16.5364 | -34.9619 | -26.2745
16.5425 | -34.9624 | -26.2822
16.5487 | -34.9629 | -26.2898
16.5548 | -34.9634 | -26.2975
16.5609 | -34.964 | -26.3051
16.567 | -34.9645 | -26.3125
16.5731 | -34.965 | -26.3201
16.5792 | -34.9655 | -26.3278
16.5853 | -34.9661 | -26.3353
16.5914 | -34.9666 | -26.3431
16.5975 | -34.9671 | -26.3508
16.6036 | -34.9677 | -26.3583
16.6097 | -34.9682 | -26.3658
16.6158 | -34.9687 | -26.3735
16.6219 | -34.9693 | -26.381
16.628 | -34.9698 | -26.3887
16.6341 | -34.9703 | -26.3962
16.6402 | -34.9709 | -26.4034
16.6463 | -34.9714 | -26.4109
16.6524 | -34.9719 | -26.4182
16.6585 | -34.9724 | -26.4259
16.6646 | -34.973 | -26.4335
16.6707 | -34.9735 | -26.4409
16.6768 | -34.974 | -26.4484
16.6829 | -34.9746 | -26.4558
16.689 | -34.9751 | -26.4635
16.6952 | -34.9756 | -26.471
16.7013 | -34.9761 | -26.4786
16.7074 | -34.9767 | -26.4864
16.7135 | -34.9772 | -26.4939
16.7196 | -34.9777 | -26.5014
16.7257 | -34.9782 | -26.5088
16.7318 | -34.9787 | -26.5163
16.7379 | -34.9793 | -26.524
16.744 | -34.9798 | -26.5315
16.7501 | -34.9803 | -26.539
16.7562 | -34.9809 | -26.5467
16.7623 | -34.9814 | -26.5542
16.7684 | -34.9819 | -26.5618
16.7745 | -34.9825 | -26.5691
16.7806 | -34.983 | -26.5768
16.7867 | -34.9835 | -26.5843
16.7928 | -34.9841 | -26.5918
16.7989 | -34.9845 | -26.5991
16.805 | -34.985 | -26.6066
16.8111 | -34.9856 | -26.6143
16.8172 | -34.9861 | -26.6222
16.8233 | -34.9866 | -26.63
16.8294 | -34.9872 | -26.6376
16.8356 | -34.9877 | -26.6453
16.8417 | -34.9882 | -26.6527
16.8478 | -34.9887 | -26.66
16.8539 | -34.9893 | -26.6675
16.86 | -34.9898 | -26.675
16.8661 | -34.9904 | -26.6825
16.8722 | -34.9909 | -26.6903
16.8783 | -34.9915 | -26.6977
16.8844 | -34.992 | -26.705
16.8905 | -34.9925 | -26.7127
16.8966 | -34.993 | -26.7204
16.9027 | -34.9936 | -26.7278
16.9088 | -34.9941 | -26.7357
16.9149 | -34.9947 | -26.7434
16.921 | -34.9952 | -26.7509
16.9271 | -34.9958 | -26.7587
16.9332 | -34.9963 | -26.7663
16.9393 | -34.9968 | -26.7742
16.9454 | -34.9973 | -26.7817
16.9515 | -34.9978 | -26.7894
16.9576 | -34.9983 | -26.7969
16.9637 | -34.9989 | -26.8044
16.9698 | -34.9994 | -26.8116
16.9759 | -34.9999 | -26.819
16.9821 | -35.0005 | -26.8268
16.9882 | -35.001 | -26.8344
16.9943 | -35.0015 | -26.8422
17.0004 | -35.0021 | -26.8497
17.0065 | -35.0026 | -26.8572
17.0126 | -35.0031 | -26.8647
17.0187 | -35.0036 | -26.8723
17.0248 | -35.0042 | -26.8796
17.0309 | -35.0047 | -26.8872
17.037 | -35.0052 | -26.8949
17.0431 | -35.0058 | -26.9022
17.0492 | -35.0063 | -26.9098
17.0553 | -35.0068 | -26.9173
17.0614 | -35.0073 | -26.9247
17.0675 | -35.0078 | -26.9324
17.0736 | -35.0083 | -26.9399
17.0797 | -35.0089 | -26.9472
17.0858 | -35.0094 | -26.955
17.0919 | -35.0099 | -26.9627
17.098 | -35.0105 | -26.9703
17.1041 | -35.011 | -26.9776
17.1102 | -35.0115 | -26.9853
17.1163 | -35.0121 | -26.9928
17.1225 | -35.0126 | -27.0001
17.1286 | -35.0132 | -27.0074
17.1347 | -35.0137 | -27.0149
17.1408 | -35.0142 | -27.0225
17.1469 | -35.0148 | -27.0299
17.153 | -35.0153 | -27.0373
17.1591 | -35.0158 | -27.0446
17.1652 | -35.0164 | -27.0519
17.1713 | -35.0169 | -27.0594
17.1774 | -35.0174 | -27.067
17.1835 | -35.018 | -27.0745
17.1896 | -35.0185 | -27.0819
17.1957 | -35.019 | -27.0895
17.2018 | -35.0195 | -27.097
17.2079 | -35.02 | -27.1043
17.214 | -35.0206 | -27.1118
17.2201 | -35.0211 | -27.1194
17.2262 | -35.0216 | -27.1269
17.2323 | -35.0221 | -27.1344
17.2384 | -35.0226 | -27.1418
17.2445 | -35.0231 | -27.1492
17.2506 | -35.0236 | -27.1568
17.2567 | -35.0241 | -27.1644
17.2628 | -35.0247 | -27.172
17.269 | -35.0252 | -27.1796
17.2751 | -35.0257 | -27.1871
17.2812 | -35.0262 | -27.1947
17.2873 | -35.0267 | -27.2023
17.2934 | -35.0272 | -27.2097
17.2995 | -35.0277 | -27.2171
17.3056 | -35.0283 | -27.2246
17.3117 | -35.0288 | -27.232
17.3178 | -35.0293 | -27.2395
17.3239 | -35.0298 | -27.247
17.33 | -35.0303 | -27.2543
17.3361 | -35.0308 | -27.2618
17.3422 | -35.0313 | -27.2692
17.3483 | -35.0319 | -27.2768
17.3544 | -35.0324 | -27.2841
17.3605 | -35.0329 | -27.2916
17.3666 | -35.0335 | -27.2992
17.3727 | -35.034 | -27.3067
17.3788 | -35.0345 | -27.3139
17.3849 | -35.035 | -27.3212
17.391 | -35.0356 | -27.3287
17.3971 | -35.0361 | -27.3362
17.4032 | -35.0366 | -27.344
17.4094 | -35.0372 | -27.3514
17.4155 | -35.0377 | -27.3587
17.4216 | -35.0381 | -27.3661
17.4277 | -35.0387 | -27.3737
17.4338 | -35.0392 | -27.3812
17.4399 | -35.0398 | -27.3887
17.446 | -35.0403 | -27.3961
17.4521 | -35.0408 | -27.4034
17.4582 | -35.0413 | -27.4108
17.4643 | -35.0419 | -27.4183
17.4704 | -35.0424 | -27.4257
17.4765 | -35.0429 | -27.4334
17.4826 | -35.0434 | -27.4408
17.4887 | -35.0439 | -27.4486
17.4948 | -35.0444 | -27.456
17.5009 | -35.0449 | -27.4634
17.507 | -35.0455 | -27.4708
17.5131 | -35.046 | -27.4783
17.5192 | -35.0465 | -27.4856
17.5253 | -35.047 | -27.493
17.5314 | -35.0475 | -27.5004
17.5375 | -35.048 | -27.5079
17.5436 | -35.0485 | -27.5155
17.5497 | -35.049 | -27.523
17.5559 | -35.0495 | -27.5302
17.562 | -35.0501 | -27.5376
17.5681 | -35.0506 | -27.5451
17.5742 | -35.0511 | -27.5527
17.5803 | -35.0516 | -27.5601
17.5864 | -35.0522 | -27.5675
17.5925 | -35.0527 | -27.5753
17.5986 | -35.0533 | -27.5824
17.6047 | -35.0538 | -27.5901
17.6108 | -35.0544 | -27.5974
17.6169 | -35.0549 | -27.6048
17.623 | -35.0554 | -27.6121
17.6291 | -35.0559 | -27.6193
17.6352 | -35.0565 | -27.6267
17.6413 | -35.057 | -27.6343
17.6474 | -35.0575 | -27.6418
17.6535 | -35.058 | -27.6493
17.6596 | -35.0585 | -27.6567
17.6657 | -35.0591 | -27.664
17.6718 | -35.0596 | -27.6715
17.6779 | -35.0601 | -27.6788
17.684 | -35.0606 | -27.6864
17.6901 | -35.0612 | -27.6938
17.6963 | -35.0616 | -27.7011
17.7024 | -35.0622 | -27.7086
17.7085 | -35.0627 | -27.716
17.7146 | -35.0632 | -27.7233
17.7207 | -35.0638 | -27.7306
17.7268 | -35.0643 | -27.738
17.7329 | -35.0648 | -27.7456
17.739 | -35.0654 | -27.7531
17.7451 | -35.0659 | -27.7603
17.7512 | -35.0664 | -27.768
17.7573 | -35.0669 | -27.7753
17.7634 | -35.0675 | -27.7827
17.7695 | -35.068 | -27.7901
17.7756 | -35.0685 | -27.7975
17.7817 | -35.069 | -27.8047
17.7878 | -35.0696 | -27.8122
17.7939 | -35.0701 | -27.8196
17.8 | -35.0706 | -27.8268
17.8061 | -35.0711 | -27.8342
17.8122 | -35.0716 | -27.8417
17.8183 | -35.0722 | -27.8492
17.8244 | -35.0727 | -27.8566
17.8305 | -35.0732 | -27.864
17.8366 | -35.0738 | -27.8711
17.8428 | -35.0743 | -27.8787
17.8489 | -35.0748 | -27.8858
17.855 | -35.0754 | -27.893
17.8611 | -35.0759 | -27.9002
17.8672 | -35.0764 | -27.9075
17.8733 | -35.0769 | -27.9151
17.8794 | -35.0774 | -27.9226
17.8855 | -35.078 | -27.93
17.8916 | -35.0785 | -27.9371
17.8977 | -35.079 | -27.9445
17.9038 | -35.0795 | -27.9515
17.9099 | -35.0801 | -27.9588
17.916 | -35.0806 | -27.9662
17.9221 | -35.0811 | -27.9735
17.9282 | -35.0817 | -27.9808
17.9343 | -35.0822 | -27.9883
17.9404 | -35.0827 | -27.9955
17.9465 | -35.0833 | -28.0028
17.9526 | -35.0838 | -28.0103
17.9587 | -35.0843 | -28.0175
17.9648 | -35.0848 | -28.0248
17.9709 | -35.0854 | -28.0321
17.977 | -35.0859 | -28.0394
17.9832 | -35.0864 | -28.0468
17.9893 | -35.0869 | -28.0542
17.9954 | -35.0874 | -28.0619
18.0015 | -35.088 | -28.0691
18.0076 | -35.0884 | -28.0759
18.0137 | -35.0889 | -28.0834
18.0198 | -35.0894 | -28.0907
18.0259 | -35.0899 | -28.0983
18.032 | -35.0904 | -28.1056
18.0381 | -35.0909 | -28.113
18.0442 | -35.0915 | -28.1204
18.0503 | -35.092 | -28.1279
18.0564 | -35.0925 | -28.1351
18.0625 | -35.093 | -28.1426
18.0686 | -35.0935 | -28.15
18.0747 | -35.094 | -28.1572
18.0808 | -35.0946 | -28.1643
18.0869 | -35.0951 | -28.1717
18.093 | -35.0956 | -28.1791
18.0991 | -35.0961 | -28.1866
18.1052 | -35.0967 | -28.1942
18.1113 | -35.0972 | -28.2012
18.1174 | -35.0977 | -28.2084
18.1236 | -35.0983 | -28.2158
18.1297 | -35.0988 | -28.2231
18.1358 | -35.0993 | -28.2305
18.1419 | -35.0999 | -28.2378
18.148 | -35.1005 | -28.2452
18.1541 | -35.101 | -28.2522
18.1602 | -35.1015 | -28.2597
18.1663 | -35.1021 | -28.2671
18.1724 | -35.1026 | -28.2747
18.1785 | -35.1031 | -28.2821
18.1846 | -35.1036 | -28.2893
18.1907 | -35.1041 | -28.2969
18.1968 | -35.1046 | -28.3043
18.2029 | -35.1051 | -28.3112
18.209 | -35.1057 | -28.3187
18.2151 | -35.1062 | -28.3262
18.2212 | -35.1067 | -28.3333
18.2273 | -35.1073 | -28.3407
18.2334 | -35.1078 | -28.3479
18.2395 | -35.1083 | -28.3552
18.2456 | -35.1088 | -28.3624
18.2517 | -35.1093 | -28.3698
18.2578 | -35.1098 | -28.3768
18.2639 | -35.1104 | -28.3845
18.2701 | -35.1109 | -28.3918
18.2762 | -35.1115 | -28.3993
18.2823 | -35.112 | -28.4067
18.2884 | -35.1125 | -28.414
18.2945 | -35.1131 | -28.4209
18.3006 | -35.1136 | -28.4283
18.3067 | -35.1141 | -28.4357
18.3128 | -35.1147 | -28.4429
18.3189 | -35.1152 | -28.4504
18.325 | -35.1157 | -28.4573
18.3311 | -35.1162 | -28.4648
18.3372 | -35.1167 | -28.4719
18.3433 | -35.1172 | -28.479
18.3494 | -35.1177 | -28.4863
18.3555 | -35.1183 | -28.4933
18.3616 | -35.1188 | -28.5004
18.3677 | -35.1193 | -28.5077
18.3738 | -35.1199 | -28.5151
18.3799 | -35.1204 | -28.5223
18.386 | -35.1209 | -28.5295
18.3921 | -35.1214 | -28.5368
18.3982 | -35.122 | -28.5438
18.4043 | -35.1225 | -28.551
18.4105 | -35.123 | -28.5583
18.4166 | -35.1235 | -28.5654
18.4227 | -35.124 | -28.5727
18.4288 | -35.1245 | -28.5802
18.4349 | -35.125 | -28.587
18.441 | -35.1256 | -28.5941
18.4471 | -35.1261 | -28.6013
18.4532 | -35.1266 | -28.6086
18.4593 | -35.1271 | -28.6157
18.4654 | -35.1277 | -28.6226
18.4715 | -35.1282 | -28.6298
18.4776 | -35.1287 | -28.6372
18.4837 | -35.1293 | -28.6445
18.4898 | -35.1298 | -28.6518
18.4959 | -35.1303 | -28.6592
18.502 | -35.1309 | -28.6664
18.5081 | -35.1314 | -28.6737
18.5142 | -35.1319 | -28.6811
18.5203 | -35.1324 | -28.6881
18.5264 | -35.1329 | -28.6951
18.5325 | -35.1335 | -28.7025
18.5386 | -35.134 | -28.7098
18.5447 | -35.1346 | -28.717
18.5508 | -35.1351 | -28.724
18.557 | -35.1356 | -28.7312
18.5631 | -35.1361 | -28.7384
18.5692 | -35.1366 | -28.7455
18.5753 | -35.1372 | -28.7524
18.5814 | -35.1377 | -28.7597
18.5875 | -35.1382 | -28.7671
18.5936 | -35.1387 | -28.7747
18.5997 | -35.1393 | -28.7818
18.6058 | -35.1397 | -28.7889
18.6119 | -35.1403 | -28.7963
18.618 | -35.1408 | -28.8034
18.6241 | -35.1413 | -28.8108
18.6302 | -35.1419 | -28.8182
18.6363 | -35.1424 | -28.8252
18.6424 | -35.1429 | -28.8323
18.6485 | -35.1434 | -28.8394
18.6546 | -35.144 | -28.8464
18.6607 | -35.1445 | -28.8538
18.6668 | -35.145 | -28.861
18.6729 | -35.1455 | -28.8683
18.679 | -35.1461 | -28.8753
18.6851 | -35.1466 | -28.8824
18.6912 | -35.1472 | -28.8896
18.6974 | -35.1477 | -28.8968
18.7035 | -35.1482 | -28.9041
18.7096 | -35.1488 | -28.9114
18.7157 | -35.1493 | -28.9185
18.7218 | -35.1499 | -28.9259
18.7279 | -35.1504 | -28.9329
18.734 | -35.1509 | -28.9401
18.7401 | -35.1514 | -28.947
18.7462 | -35.152 | -28.9538
18.7523 | -35.1525 | -28.9609
18.7584 | -35.1531 | -28.9683
18.7645 | -35.1537 | -28.9752
18.7706 | -35.1542 | -28.9824
18.7767 | -35.1547 | -28.9895
18.7828 | -35.1552 | -28.9964
18.7889 | -35.1557 | -29.0035
18.795 | -35.1562 | -29.0106
18.8011 | -35.1568 | -29.0177
18.8072 | -35.1573 | -29.0248
18.8133 | -35.1579 | -29.032
18.8194 | -35.1584 | -29.0391
18.8255 | -35.1589 | -29.0464
18.8316 | -35.1594 | -29.0536
18.8377 | -35.1599 | -29.0606
18.8439 | -35.1604 | -29.0677
18.85 | -35.161 | -29.0748
18.8561 | -35.1615 | -29.0818
18.8622 | -35.162 | -29.0891
18.8683 | -35.1625 | -29.096
18.8744 | -35.163 | -29.1033
18.8805 | -35.1636 | -29.1105
18.8866 | -35.1641 | -29.1179
18.8927 | -35.1646 | -29.125
18.8988 | -35.1652 | -29.1321
18.9049 | -35.1657 | -29.1392
18.911 | -35.1662 | -29.1463
18.9171 | -35.1668 | -29.1534
18.9232 | -35.1673 | -29.1605
18.9293 | -35.1678 | -29.1676
18.9354 | -35.1683 | -29.1744
18.9415 | -35.1688 | -29.1814
18.9476 | -35.1693 | -29.1887
18.9537 | -35.1698 | -29.1959
18.9598 | -35.1704 | -29.2024
18.9659 | -35.171 | -29.2094
18.972 | -35.1714 | -29.2165
18.9781 | -35.172 | -29.2236
18.9843 | -35.1725 | -29.2307
18.9904 | -35.173 | -29.2379
18.9965 | -35.1735 | -29.245
19.0026 | -35.1741 | -29.2525
19.0087 | -35.1746 | -29.2598
19.0148 | -35.1752 | -29.2669
19.0209 | -35.1757 | -29.2743
19.027 | -35.1762 | -29.2816
19.0331 | -35.1768 | -29.2887
19.0392 | -35.1773 | -29.2959
19.0453 | -35.1779 | -29.303
19.0514 | -35.1783 | -29.3101
19.0575 | -35.1789 | -29.3173
19.0636 | -35.1794 | -29.324
19.0697 | -35.18 | -29.3312
19.0758 | -35.1805 | -29.3382
19.0819 | -35.181 | -29.3455
19.088 | -35.1815 | -29.3523
19.0941 | -35.182 | -29.3592
19.1002 | -35.1825 | -29.366
19.1063 | -35.1831 | -29.3733
19.1124 | -35.1837 | -29.3804
19.1185 | -35.1842 | -29.3875
19.1246 | -35.1847 | -29.3944
19.1308 | -35.1852 | -29.4012
19.1369 | -35.1858 | -29.4081
19.143 | -35.1863 | -29.4151
19.1491 | -35.1868 | -29.4221
19.1552 | -35.1874 | -29.4293
19.1613 | -35.1879 | -29.4363
19.1674 | -35.1884 | -29.4434
19.1735 | -35.189 | -29.4506
19.1796 | -35.1895 | -29.4575
19.1857 | -35.19 | -29.4649
19.1918 | -35.1905 | -29.4722
19.1979 | -35.1911 | -29.4791
19.204 | -35.1916 | -29.486
19.2101 | -35.1921 | -29.4931
19.2162 | -35.1926 | -29.5003
19.2223 | -35.1931 | -29.507
19.2284 | -35.1937 | -29.5143
19.2345 | -35.1942 | -29.5213
19.2406 | -35.1948 | -29.5282
19.2467 | -35.1953 | -29.5353
19.2528 | -35.1958 | -29.5424
19.2589 | -35.1963 | -29.5493
19.265 | -35.1968 | -29.5564
19.2712 | -35.1973 | -29.5635
19.2773 | -35.1978 | -29.5704
19.2834 | -35.1983 | -29.5773
19.2895 | -35.1989 | -29.5844
19.2956 | -35.1995 | -29.5916
19.3017 | -35.2 | -29.5984
19.3078 | -35.2005 | -29.6054
19.3139 | -35.201 | -29.6124
19.32 | -35.2016 | -29.6194
19.3261 | -35.202 | -29.6262
19.3322 | -35.2026 | -29.6331
19.3383 | -35.2031 | -29.6401
19.3444 | -35.2037 | -29.6469
19.3505 | -35.2042 | -29.6537
19.3566 | -35.2046 | -29.6608
19.3627 | -35.2052 | -29.6678
19.3688 | -35.2057 | -29.675
19.3749 | -35.2063 | -29.6823
19.381 | -35.2069 | -29.6889
19.3871 | -35.2073 | -29.6958
19.3932 | -35.2079 | -29.7032
19.3993 | -35.2084 | -29.7101
19.4054 | -35.2089 | -29.7171
19.4115 | -35.2095 | -29.7242
19.4177 | -35.2099 | -29.7306
19.4238 | -35.2105 | -29.7372
19.4299 | -35.211 | -29.7441
19.436 | -35.2116 | -29.7514
19.4421 | -35.2121 | -29.758
19.4482 | -35.2127 | -29.7652
19.4543 | -35.2132 | -29.772
19.4604 | -35.2137 | -29.7789
19.4665 | -35.2142 | -29.7857
19.4726 | -35.2148 | -29.7925
19.4787 | -35.2153 | -29.7995
19.4848 | -35.2158 | -29.8064
19.4909 | -35.2163 | -29.8133
19.497 | -35.2168 | -29.82
19.5031 | -35.2173 | -29.827
19.5092 | -35.2178 | -29.8337
19.5153 | -35.2183 | -29.841
19.5214 | -35.2189 | -29.8478
19.5275 | -35.2194 | -29.8547
19.5336 | -35.2199 | -29.8616
19.5397 | -35.2205 | -29.8686
19.5458 | -35.221 | -29.8754
19.5519 | -35.2215 | -29.8821
19.5581 | -35.222 | -29.8888
19.5642 | -35.2225 | -29.8957
19.5703 | -35.2231 | -29.9025
19.5764 | -35.2236 | -29.9095
19.5825 | -35.2241 | -29.9164
19.5886 | -35.2247 | -29.9232
19.5947 | -35.2252 | -29.9298
19.6008 | -35.2257 | -29.9367
19.6069 | -35.2262 | -29.9433
19.613 | -35.2268 | -29.9503
19.6191 | -35.2273 | -29.9573
19.6252 | -35.2278 | -29.9643
19.6313 | -35.2283 | -29.9707
19.6374 | -35.2288 | -29.9776
19.6435 | -35.2293 | -29.9845
19.6496 | -35.2299 | -29.9913
19.6557 | -35.2304 | -29.9981
19.6618 | -35.2309 | -30.005
19.6679 | -35.2314 | -30.0119
19.674 | -35.232 | -30.0187
19.6801 | -35.2325 | -30.0255
19.6862 | -35.233 | -30.0325
19.6923 | -35.2335 | -30.0391
19.6984 | -35.234 | -30.0456
19.7046 | -35.2345 | -30.0523
19.7107 | -35.2351 | -30.0596
19.7168 | -35.2355 | -30.0667
19.7229 | -35.236 | -30.0736
19.729 | -35.2365 | -30.0804
19.7351 | -35.237 | -30.0866
19.7412 | -35.2375 | -30.0934
19.7473 | -35.238 | -30.1004
19.7534 | -35.2385 | -30.1074
19.7595 | -35.239 | -30.1141
19.7656 | -35.2395 | -30.1213
19.7717 | -35.24 | -30.1278
19.7778 | -35.2405 | -30.135
19.7839 | -35.241 | -30.1416
19.79 | -35.2415 | -30.1489
19.7961 | -35.242 | -30.1559
19.8022 | -35.2425 | -30.1623
19.8083 | -35.2431 | -30.1692
19.8144 | -35.2436 | -30.1758
19.8205 | -35.2441 | -30.1827
19.8266 | -35.2446 | -30.1895
19.8327 | -35.2451 | -30.1967
19.8388 | -35.2456 | -30.2034
19.845 | -35.2461 | -30.2102
19.8511 | -35.2466 | -30.2174
19.8572 | -35.2471 | -30.2242
19.8633 | -35.2476 | -30.2312
19.8694 | -35.2481 | -30.2381
19.8755 | -35.2487 | -30.2449
19.8816 | -35.2492 | -30.252
19.8877 | -35.2497 | -30.2589
19.8938 | -35.2502 | -30.2661
19.8999 | -35.2507 | -30.2728
19.906 | -35.2512 | -30.2798
19.9121 | -35.2517 | -30.2862
19.9182 | -35.2522 | -30.2929
19.9243 | -35.2527 | -30.2997
19.9304 | -35.2532 | -30.3066
19.9365 | -35.2537 | -30.3137
19.9426 | -35.2542 | -30.3202
19.9487 | -35.2548 | -30.3269
19.9548 | -35.2553 | -30.3341
19.9609 | -35.2558 | -30.3407
19.967 | -35.2564 | -30.3475
19.9731 | -35.2569 | -30.3543
19.9792 | -35.2573 | -30.3609
19.9853 | -35.2579 | -30.3679
19.9915 | -35.2584 | -30.3748
19.9976 | -35.2589 | -30.3815
20.0037 | -35.2594 | -30.3881
20.0098 | -35.2599 | -30.3948
20.0159 | -35.2604 | -30.4017
20.022 | -35.2609 | -30.4087
20.0281 | -35.2615 | -30.4158
20.0342 | -35.262 | -30.4225
20.0403 | -35.2625 | -30.4291
20.0464 | -35.263 | -30.4362
20.0525 | -35.2635 | -30.4424
20.0586 | -35.264 | -30.4491
20.0647 | -35.2646 | -30.4562
20.0708 | -35.2651 | -30.4631
20.0769 | -35.2656 | -30.4702
20.083 | -35.2661 | -30.4769
20.0891 | -35.2666 | -30.4837
20.0952 | -35.2671 | -30.4902
20.1013 | -35.2676 | -30.4971
20.1074 | -35.2681 | -30.504
20.1135 | -35.2687 | -30.5106
20.1196 | -35.2691 | -30.517
20.1257 | -35.2696 | -30.5238
20.1319 | -35.2702 | -30.5307
20.138 | -35.2707 | -30.5373
20.1441 | -35.2712 | -30.5438
20.1502 | -35.2717 | -30.5509
20.1563 | -35.2722 | -30.5574
20.1624 | -35.2727 | -30.5641
20.1685 | -35.2732 | -30.5711
20.1746 | -35.2737 | -30.5781
20.1807 | -35.2742 | -30.5851
20.1868 | -35.2747 | -30.5918
20.1929 | -35.2753 | -30.5983
20.199 | -35.2757 | -30.6049
20.2051 | -35.2763 | -30.6116
20.2112 | -35.2768 | -30.6182
20.2173 | -35.2773 | -30.6252
20.2234 | -35.2778 | -30.6319
20.2295 | -35.2783 | -30.6385
20.2356 | -35.2788 | -30.6453
20.2417 | -35.2793 | -30.6521
20.2478 | -35.2798 | -30.6588
20.2539 | -35.2803 | -30.6658
20.26 | -35.2808 | -30.6724
20.2661 | -35.2813 | -30.6792
20.2723 | -35.2818 | -30.6856
20.2784 | -35.2824 | -30.6926
20.2845 | -35.2829 | -30.6993
20.2906 | -35.2834 | -30.706
20.2967 | -35.2839 | -30.7126
20.3028 | -35.2844 | -30.7193
20.3089 | -35.2849 | -30.726
20.315 | -35.2854 | -30.7324
20.3211 | -35.2859 | -30.739
20.3272 | -35.2865 | -30.7458
20.3333 | -35.2869 | -30.7523
20.3394 | -35.2875 | -30.759
20.3455 | -35.2881 | -30.7656
20.3516 | -35.2886 | -30.7725
20.3577 | -35.2891 | -30.7789
20.3638 | -35.2897 | -30.7858
20.3699 | -35.2902 | -30.7919
20.376 | -35.2907 | -30.7988
20.3821 | -35.2912 | -30.8053
20.3882 | -35.2917 | -30.8122
20.3943 | -35.2922 | -30.8188
20.4004 | -35.2928 | -30.8253
20.4065 | -35.2932 | -30.832
20.4126 | -35.2938 | -30.8387
20.4188 | -35.2943 | -30.8457
20.4249 | -35.2948 | -30.8525
20.431 | -35.2952 | -30.8592
20.4371 | -35.2957 | -30.8655
20.4432 | -35.2962 | -30.8722
20.4493 | -35.2967 | -30.879
20.4554 | -35.2972 | -30.8858
20.4615 | -35.2977 | -30.8928
20.4676 | -35.2982 | -30.8994
20.4737 | -35.2987 | -30.906
20.4798 | -35.2992 | -30.9128
20.4859 | -35.2997 | -30.9194
20.492 | -35.3002 | -30.926
20.4981 | -35.3007 | -30.9327
20.5042 | -35.3012 | -30.9395
20.5103 | -35.3017 | -30.9459
20.5164 | -35.3022 | -30.9526
20.5225 | -35.3026 | -30.9596
20.5286 | -35.3031 | -30.9663
20.5347 | -35.3036 | -30.9729
20.5408 | -35.3041 | -30.9799
20.5469 | -35.3047 | -30.9863
20.553 | -35.3052 | -30.9929
20.5592 | -35.3056 | -30.9995
20.5653 | -35.306 | -31.0063
20.5714 | -35.3065 | -31.0123
20.5775 | -35.307 | -31.0192
20.5836 | -35.3075 | -31.0256
20.5897 | -35.3081 | -31.0327
20.5958 | -35.3085 | -31.0393
20.6019 | -35.309 | -31.0459
20.608 | -35.3095 | -31.0524
20.6141 | -35.3101 | -31.0589
20.6202 | -35.3106 | -31.0655
20.6263 | -35.3111 | -31.0718
20.6324 | -35.3116 | -31.0782
20.6385 | -35.3121 | -31.0852
20.6446 | -35.3126 | -31.0923
20.6507 | -35.3131 | -31.0985
20.6568 | -35.3136 | -31.1052
20.6629 | -35.3141 | -31.1121
20.669 | -35.3146 | -31.1186
20.6751 | -35.3151 | -31.1249
20.6812 | -35.3156 | -31.1309
20.6873 | -35.3162 | -31.1374
20.6934 | -35.3167 | -31.1438
20.6995 | -35.3171 | -31.1505
20.7057 | -35.3176 | -31.1571
20.7118 | -35.3181 | -31.1637
20.7179 | -35.3185 | -31.1696
20.724 | -35.319 | -31.1763
20.7301 | -35.3195 | -31.1837
20.7362 | -35.32 | -31.1905
20.7423 | -35.3205 | -31.1969
20.7484 | -35.321 | -31.2036
20.7545 | -35.3215 | -31.2099
20.7606 | -35.322 | -31.2166
20.7667 | -35.3225 | -31.2229
20.7728 | -35.3231 | -31.2294
20.7789 | -35.3236 | -31.2363
20.785 | -35.3241 | -31.243
20.7911 | -35.3246 | -31.2492
20.7972 | -35.3251 | -31.2561
20.8033 | -35.3256 | -31.2626
20.8094 | -35.3261 | -31.2691
20.8155 | -35.3266 | -31.276
20.8216 | -35.327 | -31.2829
20.8277 | -35.3275 | -31.2895
20.8338 | -35.328 | -31.2961
20.8399 | -35.3286 | -31.3022
20.8461 | -35.329 | -31.3085
20.8522 | -35.3296 | -31.3148
20.8583 | -35.3302 | -31.3213
20.8644 | -35.3307 | -31.3283
20.8705 | -35.3312 | -31.3352
20.8766 | -35.3317 | -31.3415
20.8827 | -35.3323 | -31.3474
20.8888 | -35.3328 | -31.3536
20.8949 | -35.3332 | -31.3603
20.901 | -35.3338 | -31.367
20.9071 | -35.3343 | -31.3737
20.9132 | -35.3347 | -31.38
20.9193 | -35.3352 | -31.3865
20.9254 | -35.3357 | -31.3924
20.9315 | -35.3361 | -31.3987
20.9376 | -35.3366 | -31.4055
20.9437 | -35.3372 | -31.4122
20.9498 | -35.3377 | -31.4184
20.9559 | -35.3382 | -31.4243
20.962 | -35.3387 | -31.4314
20.9681 | -35.3392 | -31.4382
20.9742 | -35.3396 | -31.4448
20.9803 | -35.3401 | -31.4513
20.9864 | -35.3405 | -31.4579
20.9926 | -35.3411 | -31.4646
20.9987 | -35.3415 | -31.4708
21.0048 | -35.3419 | -31.4775
21.0109 | -35.3425 | -31.4839
21.017 | -35.343 | -31.4904
21.0231 | -35.3435 | -31.4969
21.0292 | -35.344 | -31.5036
21.0353 | -35.3444 | -31.5096
21.0414 | -35.3449 | -31.5154
21.0475 | -35.3455 | -31.5214
21.0536 | -35.3461 | -31.5279
21.0597 | -35.3466 | -31.534
21.0658 | -35.3471 | -31.5407
21.0719 | -35.3476 | -31.547
21.078 | -35.3482 | -31.5536
21.0841 | -35.3488 | -31.5595
21.0902 | -35.3493 | -31.5654
21.0963 | -35.3497 | -31.5714
21.1024 | -35.3501 | -31.5777
21.1085 | -35.3505 | -31.5838
21.1146 | -35.351 | -31.59
21.1207 | -35.3514 | -31.5967
21.1268 | -35.3519 | -31.6033
21.133 | -35.3524 | -31.6096
21.1391 | -35.3528 | -31.6163
21.1452 | -35.3532 | -31.6226
21.1513 | -35.3537 | -31.629
21.1574 | -35.3542 | -31.6357
21.1635 | -35.3547 | -31.6419
21.1696 | -35.3552 | -31.6472
21.1757 | -35.3557 | -31.6528
21.1818 | -35.3562 | -31.6594
21.1879 | -35.3568 | -31.6661
21.194 | -35.3573 | -31.6726
21.2001 | -35.3578 | -31.6788
21.2062 | -35.3583 | -31.685
21.2123 | -35.3589 | -31.6915
21.2184 | -35.3594 | -31.6976
21.2245 | -35.3599 | -31.7043
21.2306 | -35.3604 | -31.7104
21.2367 | -35.3609 | -31.717
21.2428 | -35.3614 | -31.7225
21.2489 | -35.3618 | -31.7279
21.255 | -35.3622 | -31.7337
21.2611 | -35.3627 | -31.7399
21.2672 | -35.3632 | -31.7459
21.2733 | -35.3637 | -31.7523
21.2795 | -35.3642 | -31.7576
21.2856 | -35.3647 | -31.7645
21.2917 | -35.3652 | -31.7708
21.2978 | -35.3656 | -31.7771
21.3039 | -35.3661 | -31.7831
21.31 | -35.3665 | -31.7889
21.3161 | -35.367 | -31.7948
21.3222 | -35.3676 | -31.8003
21.3283 | -35.3681 | -31.807
21.3344 | -35.3685 | -31.8127
21.3405 | -35.3689 | -31.8195
21.3466 | -35.3695 | -31.8252
21.3527 | -35.3699 | -31.8316
21.3588 | -35.3705 | -31.8376
21.3649 | -35.371 | -31.8435
21.371 | -35.3715 | -31.8504
21.3771 | -35.372 | -31.8562
21.3832 | -35.3725 | -31.8621
21.3893 | -35.3729 | -31.8679
21.3954 | -35.3734 | -31.8741
21.4015 | -35.3739 | -31.8805
21.4076 | -35.3744 | -31.8867
21.4137 | -35.3748 | -31.8932
21.4199 | -35.3752 | -31.8988
21.426 | -35.3756 | -31.9056
21.4321 | -35.3761 | -31.9122
21.4382 | -35.3765 | -31.9178
21.4443 | -35.3769 | -31.9238
21.4504 | -35.3775 | -31.9305
21.4565 | -35.3779 | -31.9356
21.4626 | -35.3784 | -31.9418
21.4687 | -35.379 | -31.9478
21.4748 | -35.3795 | -31.9537
21.4809 | -35.38 | -31.9599
21.487 | -35.3803 | -31.9654
21.4931 | -35.381 | -31.9713
21.4992 | -35.3815 | -31.9772
21.5053 | -35.3819 | -31.983
21.5114 | -35.3824 | -31.9893
21.5175 | -35.3828 | -31.9943
21.5236 | -35.3834 | -32.0006
21.5297 | -35.384 | -32.0057
21.5358 | -35.3844 | -32.011
21.5419 | -35.385 | -32.0175
21.548 | -35.3855 | -32.0239
21.5541 | -35.3861 | -32.0304
21.5602 | -35.3864 | -32.0375
21.5664 | -35.387 | -32.0433
21.5725 | -35.3874 | -32.0497
21.5786 | -35.3878 | -32.0555
21.5847 | -35.3883 | -32.0614
21.5908 | -35.3888 | -32.0674
21.5969 | -35.3892 | -32.0733
21.603 | -35.3898 | -32.0792
21.6091 | -35.3903 | -32.0855
21.6152 | -35.3907 | -32.0913
21.6213 | -35.3912 | -32.0975
21.6274 | -35.3916 | -32.1041
21.6335 | -35.3921 | -32.1098
21.6396 | -35.3926 | -32.1166
21.6457 | -35.3931 | -32.1227
21.6518 | -35.3937 | -32.1293
21.6579 | -35.3942 | -32.1351
21.664 | -35.3948 | -32.142
21.6701 | -35.3952 | -32.1479
21.6762 | -35.3957 | -32.154
21.6823 | -35.3962 | -32.1603
21.6884 | -35.3966 | -32.167
21.6945 | -35.3971 | -32.1733
21.7006 | -35.3975 | -32.1798
21.7068 | -35.3981 | -32.1849
21.7129 | -35.3985 | -32.1909
21.719 | -35.3991 | -32.1975
21.7251 | -35.3996 | -32.2033
21.7312 | -35.4001 | -32.2087
21.7373 | -35.4005 | -32.2144
21.7434 | -35.401 | -32.2208
21.7495 | -35.4015 | -32.2264
21.7556 | -35.4019 | -32.233
21.7617 | -35.4024 | -32.2391
21.7678 | -35.4029 | -32.2451
21.7739 | -35.4034 | -32.251
21.78 | -35.4039 | -32.2565
21.7861 | -35.4044 | -32.2623
21.7922 | -35.4049 | -32.2682
21.7983 | -35.4054 | -32.2741
21.8044 | -35.406 | -32.2806
21.8105 | -35.4066 | -32.2859
21.8166 | -35.4071 | -32.2914
21.8227 | -35.4076 | -32.2976
21.8288 | -35.4081 | -32.3036
21.8349 | -35.4085 | -32.3092
21.841 | -35.409 | -32.3156
21.8471 | -35.4095 | -32.3211
21.8533 | -35.41 | -32.3266
21.8594 | -35.4105 | -32.3337
21.8655 | -35.411 | -32.3397
21.8716 | -35.4116 | -32.3465
21.8777 | -35.4119 | -32.3526
21.8838 | -35.4125 | -32.3585
21.8899 | -35.413 | -32.3645
21.896 | -35.4135 | -32.3712
21.9021 | -35.4141 | -32.3764
21.9082 | -35.4145 | -32.3826
21.9143 | -35.4151 | -32.3882
21.9204 | -35.4156 | -32.3943
21.9265 | -35.4162 | -32.4001
21.9326 | -35.4168 | -32.4056
21.9387 | -35.4174 | -32.412
21.9448 | -35.4179 | -32.4176
21.9509 | -35.4185 | -32.4232
21.957 | -35.419 | -32.4294
21.9631 | -35.4195 | -32.4358
21.9692 | -35.42 | -32.4419
21.9753 | -35.4205 | -32.4473
21.9814 | -35.4211 | -32.454
21.9875 | -35.4214 | -32.4593
21.9937 | -35.422 | -32.4655
21.9998 | -35.4225 | -32.4718
22.0059 | -35.423 | -32.4776
22.012 | -35.4235 | -32.4834
22.0181 | -35.424 | -32.4899
22.0242 | -35.4244 | -32.4953
22.0303 | -35.4248 | -32.5012
22.0364 | -35.4253 | -32.5075
22.0425 | -35.4258 | -32.5132
22.0486 | -35.4264 | -32.5198
22.0547 | -35.4269 | -32.5255
22.0608 | -35.4273 | -32.5309
22.0669 | -35.4277 | -32.5374
22.073 | -35.4282 | -32.543
22.0791 | -35.4287 | -32.548
22.0852 | -35.4292 | -32.5538
22.0913 | -35.4297 | -32.5592
22.0974 | -35.4303 | -32.5653
22.1035 | -35.4306 | -32.5713
22.1096 | -35.4312 | -32.5777
22.1157 | -35.4316 | -32.5837
22.1218 | -35.4321 | -32.5899
22.1279 | -35.4327 | -32.5951
22.134 | -35.4332 | -32.6013
22.1402 | -35.4336 | -32.6071
22.1463 | -35.4341 | -32.6128
22.1524 | -35.4346 | -32.6193
22.1585 | -35.4351 | -32.6255
22.1646 | -35.4357 | -32.6313
22.1707 | -35.4361 | -32.637
22.1768 | -35.4366 | -32.6431
22.1829 | -35.4371 | -32.6487
22.189 | -35.4376 | -32.6553
22.1951 | -35.4382 | -32.6606
22.2012 | -35.4387 | -32.6667
22.2073 | -35.4391 | -32.6726
22.2134 | -35.4396 | -32.6787
22.2195 | -35.4401 | -32.6842
22.2256 | -35.4406 | -32.69
22.2317 | -35.441 | -32.6958
22.2378 | -35.4415 | -32.7014
22.2439 | -35.442 | -32.7076
22.25 | -35.4424 | -32.7136
22.2561 | -35.443 | -32.72
22.2622 | -35.4435 | -32.7261
22.2683 | -35.4439 | -32.7305
22.2744 | -35.4444 | -32.7361
22.2806 | -35.4449 | -32.7418
22.2867 | -35.4454 | -32.7479
22.2928 | -35.4458 | -32.754
22.2989 | -35.4463 | -32.7604
22.305 | -35.4468 | -32.7652
22.3111 | -35.4473 | -32.7708
22.3172 | -35.4479 | -32.777
22.3233 | -35.4483 | -32.7826
22.3294 | -35.4488 | -32.7884
22.3355 | -35.4493 | -32.7936
22.3416 | -35.4499 | -32.7992
22.3477 | -35.4504 | -32.8053
22.3538 | -35.4509 | -32.811
22.3599 | -35.4514 | -32.817
22.366 | -35.4518 | -32.8228
22.3721 | -35.4524 | -32.8279
22.3782 | -35.4529 | -32.8341
22.3843 | -35.4535 | -32.8401
22.3904 | -35.454 | -32.8462
22.3965 | -35.4545 | -32.8517
22.4026 | -35.455 | -32.8578
22.4087 | -35.4553 | -32.8636
22.4148 | -35.4558 | -32.8689
22.4209 | -35.4562 | -32.875
22.4271 | -35.4567 | -32.8811
22.4332 | -35.4572 | -32.8864
22.4393 | -35.4578 | -32.8926
22.4454 | -35.4583 | -32.8969
22.4515 | -35.4588 | -32.9028
22.4576 | -35.4592 | -32.9083
22.4637 | -35.4597 | -32.9141
22.4698 | -35.4601 | -32.9199
22.4759 | -35.4605 | -32.9249
22.482 | -35.4609 | -32.931
22.4881 | -35.4614 | -32.9371
22.4942 | -35.462 | -32.9434
22.5003 | -35.4625 | -32.9492
22.5064 | -35.463 | -32.9549
22.5125 | -35.4635 | -32.9611
22.5186 | -35.464 | -32.9674
22.5247 | -35.4644 | -32.9733
22.5308 | -35.4648 | -32.9792
22.5369 | -35.4653 | -32.9857
22.543 | -35.4658 | -32.9917
22.5491 | -35.4662 | -32.9964
22.5552 | -35.4666 | -33.0019
22.5613 | -35.467 | -33.0081
22.5675 | -35.4675 | -33.0135
22.5736 | -35.4679 | -33.0193
22.5797 | -35.4683 | -33.0253
22.5858 | -35.4688 | -33.0313
22.5919 | -35.4692 | -33.0374
22.598 | -35.4698 | -33.0425
22.6041 | -35.4703 | -33.0474
22.6102 | -35.4709 | -33.0529
22.6163 | -35.4716 | -33.0596
22.6224 | -35.4721 | -33.0654
22.6285 | -35.4726 | -33.0707
22.6346 | -35.4731 | -33.0758
22.6407 | -35.4735 | -33.0809
22.6468 | -35.4741 | -33.0866
22.6529 | -35.4747 | -33.0928
22.659 | -35.4751 | -33.0982
22.6651 | -35.4755 | -33.104
22.6712 | -35.476 | -33.1096
22.6773 | -35.4763 | -33.1149
22.6834 | -35.4768 | -33.1201
22.6895 | -35.4774 | -33.1264
22.6956 | -35.4779 | -33.1317
22.7017 | -35.4784 | -33.1372
22.7079 | -35.4789 | -33.1431
22.714 | -35.4795 | -33.1481
22.7201 | -35.48 | -33.1541
22.7262 | -35.4805 | -33.1592
22.7323 | -35.481 | -33.1648
22.7384 | -35.4816 | -33.1704
22.7445 | -35.4822 | -33.1762
22.7506 | -35.4826 | -33.1814
22.7567 | -35.4831 | -33.187
22.7628 | -35.4835 | -33.1925
22.7689 | -35.4841 | -33.1982
22.775 | -35.4847 | -33.204
22.7811 | -35.4852 | -33.2094
22.7872 | -35.4857 | -33.2149
22.7933 | -35.4861 | -33.2195
22.7994 | -35.4866 | -33.2252
22.8055 | -35.487 | -33.2306
22.8116 | -35.4875 | -33.2365
22.8177 | -35.4879 | -33.2421
22.8238 | -35.4884 | -33.2471
22.8299 | -35.489 | -33.2525
22.836 | -35.4895 | -33.2588
22.8421 | -35.49 | -33.2648
22.8482 | -35.4904 | -33.2715
22.8544 | -35.4909 | -33.2774
22.8605 | -35.4914 | -33.2831
22.8666 | -35.4919 | -33.2873
22.8727 | -35.4924 | -33.293
22.8788 | -35.4929 | -33.2983
22.8849 | -35.4934 | -33.3034
22.891 | -35.4939 | -33.3086
22.8971 | -35.4943 | -33.3142
22.9032 | -35.4948 | -33.319
22.9093 | -35.4952 | -33.3244
22.9154 | -35.4958 | -33.33
22.9215 | -35.4963 | -33.3355
22.9276 | -35.4967 | -33.3412
22.9337 | -35.4972 | -33.3473
22.9398 | -35.4977 | -33.3528
22.9459 | -35.4982 | -33.3583
22.952 | -35.4987 | -33.3639
22.9581 | -35.4992 | -33.3699
22.9642 | -35.4997 | -33.376
22.9703 | -35.5003 | -33.3814
22.9764 | -35.5008 | -33.3872
22.9825 | -35.5013 | -33.3926
22.9886 | -35.5018 | -33.3984
22.9948 | -35.5023 | -33.4037
23.0009 | -35.5027 | -33.4096
23.007 | -35.5033 | -33.4149
23.0131 | -35.5038 | -33.4207
23.0192 | -35.5043 | -33.4264
23.0253 | -35.5048 | -33.4324
23.0314 | -35.5052 | -33.4382
23.0375 | -35.5057 | -33.4446
23.0436 | -35.5063 | -33.45
23.0497 | -35.5068 | -33.4555
23.0558 | -35.5074 | -33.4607
23.0619 | -35.5078 | -33.466
23.068 | -35.5082 | -33.4711
23.0741 | -35.5088 | -33.4766
23.0802 | -35.5092 | -33.4818
23.0863 | -35.5097 | -33.4875
23.0924 | -35.5101 | -33.4927
23.0985 | -35.5105 | -33.4982
23.1046 | -35.511 | -33.5041
23.1107 | -35.5114 | -33.5086
23.1168 | -35.5119 | -33.5142
23.1229 | -35.5124 | -33.5194
23.129 | -35.5129 | -33.5249
23.1351 | -35.5134 | -33.5303
23.1413 | -35.5139 | -33.5364
23.1474 | -35.5144 | -33.5413
23.1535 | -35.5149 | -33.5471
23.1596 | -35.5154 | -33.5529
23.1657 | -35.5159 | -33.5582
23.1718 | -35.5164 | -33.564
23.1779 | -35.5168 | -33.5698
23.184 | -35.5172 | -33.5756
23.1901 | -35.5177 | -33.5811
23.1962 | -35.5181 | -33.5863
23.2023 | -35.5186 | -33.5915
23.2084 | -35.5191 | -33.5971
23.2145 | -35.5196 | -33.6027
23.2206 | -35.5201 | -33.608
23.2267 | -35.5207 | -33.6134
23.2328 | -35.5212 | -33.6189
23.2389 | -35.5216 | -33.6248
23.245 | -35.5222 | -33.6301
23.2511 | -35.5228 | -33.6361
23.2572 | -35.5234 | -33.642
23.2633 | -35.5239 | -33.6477
23.2694 | -35.5244 | -33.6538
23.2755 | -35.5249 | -33.6596
23.2817 | -35.5253 | -33.665
23.2878 | -35.5259 | -33.6708
23.2939 | -35.5264 | -33.6767
23.3 | -35.5269 | -33.6822
23.3061 | -35.5274 | -33.6877
23.3122 | -35.5279 | -33.6929
23.3183 | -35.5284 | -33.6986
23.3244 | -35.5289 | -33.7035
23.3305 | -35.5295 | -33.709
23.3366 | -35.5299 | -33.7147
23.3427 | -35.5304 | -33.7199
23.3488 | -35.5309 | -33.7259
23.3549 | -35.5315 | -33.7315
23.361 | -35.5321 | -33.7364
23.3671 | -35.5326 | -33.742
23.3732 | -35.5331 | -33.7475
23.3793 | -35.5336 | -33.7527
23.3854 | -35.5341 | -33.7582
23.3915 | -35.5346 | -33.7644
23.3976 | -35.5352 | -33.7698
23.4037 | -35.5357 | -33.7756
23.4098 | -35.5362 | -33.7807
23.4159 | -35.5366 | -33.786
23.422 | -35.5371 | -33.7911
23.4282 | -35.5378 | -33.7964
23.4343 | -35.5382 | -33.8021
23.4404 | -35.5386 | -33.8081
23.4465 | -35.5392 | -33.8137
23.4526 | -35.5397 | -33.8193
23.4587 | -35.5403 | -33.8248
23.4648 | -35.5406 | -33.829
23.4709 | -35.541 | -33.8348
23.477 | -35.5415 | -33.84
23.4831 | -35.5418 | -33.8448
23.4892 | -35.5424 | -33.8501
23.4953 | -35.5428 | -33.8557
23.5014 | -35.5434 | -33.8602
23.5075 | -35.5439 | -33.8651
23.5136 | -35.5444 | -33.8699
23.5197 | -35.5449 | -33.8757
23.5258 | -35.5454 | -33.8814
23.5319 | -35.5457 | -33.8866
23.538 | -35.5461 | -33.8922
23.5441 | -35.5466 | -33.898
23.5502 | -35.547 | -33.9033
23.5563 | -35.5475 | -33.9084
23.5624 | -35.548 | -33.9138
23.5686 | -35.5486 | -33.9191
23.5747 | -35.5491 | -33.9248
23.5808 | -35.5496 | -33.9303
23.5869 | -35.55 | -33.935
23.593 | -35.5505 | -33.94
23.5991 | -35.551 | -33.9455
23.6052 | -35.5515 | -33.9502
23.6113 | -35.552 | -33.9561
23.6174 | -35.5525 | -33.9612
23.6235 | -35.553 | -33.9663
23.6296 | -35.5534 | -33.9725
23.6357 | -35.554 | -33.9785
23.6418 | -35.5545 | -33.9835
23.6479 | -35.5549 | -33.9889
23.654 | -35.5554 | -33.9946
23.6601 | -35.5558 | -33.9996
23.6662 | -35.5563 | -34.0058
23.6723 | -35.5568 | -34.0111
23.6784 | -35.5573 | -34.0167
23.6845 | -35.5578 | -34.022
23.6906 | -35.5582 | -34.0276
23.6967 | -35.5587 | -34.0336
23.7028 | -35.5591 | -34.0388
23.7089 | -35.5597 | -34.0444
23.7151 | -35.5602 | -34.0498
23.7212 | -35.5607 | -34.055
23.7273 | -35.5612 | -34.0606
23.7334 | -35.5617 | -34.0659
23.7395 | -35.5622 | -34.0714
23.7456 | -35.5626 | -34.0775
23.7517 | -35.563 | -34.083
23.7578 | -35.5634 | -34.0886
23.7639 | -35.5639 | -34.0942
23.77 | -35.5644 | -34.1001
23.7761 | -35.5649 | -34.1053
23.7822 | -35.5652 | -34.1104
23.7883 | -35.5657 | -34.1156
23.7944 | -35.5662 | -34.1206
23.8005 | -35.5667 | -34.1258
23.8066 | -35.5672 | -34.1314
23.8127 | -35.5676 | -34.1366
23.8188 | -35.5681 | -34.1417
23.8249 | -35.5686 | -34.1473
23.831 | -35.5691 | -34.1531
23.8371 | -35.5696 | -34.1585
23.8432 | -35.5701 | -34.1638
23.8493 | -35.5705 | -34.1695
23.8555 | -35.5711 | -34.1751
23.8616 | -35.5716 | -34.1807
23.8677 | -35.5722 | -34.186
23.8738 | -35.5726 | -34.1917
23.8799 | -35.5731 | -34.197
23.886 | -35.5735 | -34.2023
23.8921 | -35.5739 | -34.2082
23.8982 | -35.5744 | -34.2137
23.9043 | -35.5749 | -34.2196
23.9104 | -35.5753 | -34.2245
23.9165 | -35.5758 | -34.2291
23.9226 | -35.5763 | -34.2343
23.9287 | -35.5767 | -34.2396
23.9348 | -35.5772 | -34.2448
23.9409 | -35.5776 | -34.2503
23.947 | -35.5782 | -34.2556
23.9531 | -35.5786 | -34.2605
23.9592 | -35.5791 | -34.2658
23.9653 | -35.5796 | -34.2711
23.9714 | -35.5801 | -34.2763
23.9775 | -35.5806 | -34.2819
23.9836 | -35.5811 | -34.2874
23.9897 | -35.5816 | -34.2932
23.9958 | -35.5821 | -34.299
24.002 | -35.5825 | -34.3036
24.0081 | -35.5829 | -34.3093
24.0142 | -35.5834 | -34.3152
24.0203 | -35.5839 | -34.3209
24.0264 | -35.5843 | -34.3264
24.0325 | -35.5849 | -34.3319
24.0386 | -35.5854 | -34.3366
24.0447 | -35.5858 | -34.342
24.0508 | -35.5863 | -34.3478
24.0569 | -35.5868 | -34.353
24.063 | -35.5871 | -34.3582
24.0691 | -35.5876 | -34.3635
24.0752 | -35.588 | -34.3688
24.0813 | -35.5884 | -34.3741
24.0874 | -35.5889 | -34.3798
24.0935 | -35.5894 | -34.3844
24.0996 | -35.5898 | -34.3898
24.1057 | -35.5903 | -34.395
24.1118 | -35.5908 | -34.4004
24.1179 | -35.5912 | -34.4063
24.124 | -35.5917 | -34.4113
24.1301 | -35.592 | -34.4168
24.1362 | -35.5924 | -34.4227
24.1424 | -35.5929 | -34.4278
24.1485 | -35.5933 | -34.4323
24.1546 | -35.5938 | -34.4376
24.1607 | -35.5944 | -34.4431
24.1668 | -35.5948 | -34.4477
24.1729 | -35.5954 | -34.4529
24.179 | -35.5958 | -34.4584
24.1851 | -35.5963 | -34.4639
24.1912 | -35.5968 | -34.4693
24.1973 | -35.5972 | -34.4744
24.2034 | -35.5976 | -34.4796
24.2095 | -35.5981 | -34.4853
24.2156 | -35.5986 | -34.491
24.2217 | -35.5991 | -34.496
24.2278 | -35.5997 | -34.5015
24.2339 | -35.6002 | -34.5073
24.24 | -35.6006 | -34.5127
24.2461 | -35.6009 | -34.5183
24.2522 | -35.6013 | -34.5242
24.2583 | -35.6017 | -34.5299
24.2644 | -35.6021 | -34.5353
24.2705 | -35.6026 | -34.5411
24.2766 | -35.603 | -34.5468
24.2827 | -35.6036 | -34.5516
24.2889 | -35.604 | -34.5568
24.295 | -35.6046 | -34.562
24.3011 | -35.605 | -34.5676
24.3072 | -35.6054 | -34.5731
24.3133 | -35.6059 | -34.5786
24.3194 | -35.6064 | -34.5843
24.3255 | -35.6069 | -34.5899
24.3316 | -35.6073 | -34.5955
24.3377 | -35.6078 | -34.6007
24.3438 | -35.6082 | -34.6066
24.3499 | -35.6087 | -34.6118
24.356 | -35.6091 | -34.6173
24.3621 | -35.6095 | -34.6229
24.3682 | -35.61 | -34.628
24.3743 | -35.6105 | -34.6332
24.3804 | -35.611 | -34.6382
24.3865 | -35.6114 | -34.644
24.3926 | -35.6118 | -34.6497
24.3987 | -35.6122 | -34.6559
24.4048 | -35.6127 | -34.6615
24.4109 | -35.6131 | -34.6667
24.417 | -35.6136 | -34.6721
24.4231 | -35.6141 | -34.6776
24.4293 | -35.6146 | -34.683
24.4354 | -35.6151 | -34.6885
24.4415 | -35.6154 | -34.6937
24.4476 | -35.6158 | -34.699
24.4537 | -35.6163 | -34.705
24.4598 | -35.6168 | -34.7107
24.4659 | -35.6172 | -34.716
24.472 | -35.6176 | -34.7218
24.4781 | -35.6181 | -34.7277
24.4842 | -35.6184 | -34.7331
24.4903 | -35.6189 | -34.7386
24.4964 | -35.6194 | -34.7443
24.5025 | -35.6199 | -34.7497
24.5086 | -35.6203 | -34.7556
24.5147 | -35.6208 | -34.761
24.5208 | -35.6213 | -34.7661
24.5269 | -35.6218 | -34.7711
24.533 | -35.6223 | -34.7766
24.5391 | -35.6228 | -34.7818
24.5452 | -35.6234 | -34.7872
24.5513 | -35.6238 | -34.793
24.5574 | -35.6243 | -34.7985
24.5635 | -35.6247 | -34.8035
24.5696 | -35.6251 | -34.8088
24.5758 | -35.6256 | -34.8139
24.5819 | -35.626 | -34.8191
24.588 | -35.6265 | -34.8248
24.5941 | -35.627 | -34.8304
24.6002 | -35.6274 | -34.836
24.6063 | -35.6279 | -34.8412
24.6124 | -35.6284 | -34.8466
24.6185 | -35.6289 | -34.8523
24.6246 | -35.6293 | -34.858
24.6307 | -35.6297 | -34.8633
24.6368 | -35.6302 | -34.8685
24.6429 | -35.6307 | -34.8739
24.649 | -35.6312 | -34.8789
24.6551 | -35.6317 | -34.8842
24.6612 | -35.6321 | -34.8891
24.6673 | -35.6326 | -34.8942
24.6734 | -35.633 | -34.9001
24.6795 | -35.6335 | -34.906
24.6856 | -35.6339 | -34.9116
24.6917 | -35.6343 | -34.9167
24.6978 | -35.6347 | -34.9223
24.7039 | -35.6351 | -34.9275
24.71 | -35.6356 | -34.9329
24.7162 | -35.636 | -34.9389
24.7223 | -35.6365 | -34.944
24.7284 | -35.6369 | -34.9504
24.7345 | -35.6373 | -34.9556
24.7406 | -35.6376 | -34.9614
24.7467 | -35.6381 | -34.9673
24.7528 | -35.6385 | -34.9739
24.7589 | -35.6389 | -34.9796
24.765 | -35.6395 | -34.9849
24.7711 | -35.64 | -34.9903
24.7772 | -35.6405 | -34.9957
24.7833 | -35.6409 | -35.0014
24.7894 | -35.6412 | -35.007
24.7955 | -35.6416 | -35.0126
24.8016 | -35.642 | -35.0185
24.8077 | -35.6425 | -35.024
24.8138 | -35.643 | -35.0297
24.8199 | -35.6434 | -35.0342
24.826 | -35.6438 | -35.0399
24.8321 | -35.6442 | -35.0457
24.8382 | -35.6448 | -35.0517
24.8443 | -35.6452 | -35.058
24.8504 | -35.6456 | -35.0633
24.8565 | -35.6461 | -35.0696
24.8627 | -35.6464 | -35.0744
24.8688 | -35.6468 | -35.0809
24.8749 | -35.6472 | -35.0859
24.881 | -35.6477 | -35.0912
24.8871 | -35.6481 | -35.0965
24.8932 | -35.6485 | -35.1018
24.8993 | -35.6489 | -35.1073
24.9054 | -35.6494 | -35.1134
24.9115 | -35.6498 | -35.1185
24.9176 | -35.6503 | -35.1242
24.9237 | -35.6508 | -35.1302
24.9298 | -35.6513 | -35.136
24.9359 | -35.6517 | -35.1413
24.942 | -35.6522 | -35.1466
24.9481 | -35.6527 | -35.1518
24.9542 | -35.6532 | -35.1572
24.9603 | -35.6536 | -35.1627
24.9664 | -35.6541 | -35.1682
24.9725 | -35.6545 | -35.1736
24.9786 | -35.6549 | -35.1792
24.9847 | -35.6553 | -35.1851
24.9908 | -35.6558 | -35.1902
24.9969 | -35.6562 | -35.196
25.0031 | -35.6565 | -35.2017
25.0092 | -35.657 | -35.2073
25.0153 | -35.6574 | -35.2125
25.0214 | -35.6578 | -35.2183
25.0275 | -35.6582 | -35.2234
25.0336 | -35.6586 | -35.2284
25.0397 | -35.659 | -35.2343
25.0458 | -35.6592 | -35.2402
25.0519 | -35.6597 | -35.2466
25.058 | -35.6602 | -35.2524
25.0641 | -35.6606 | -35.2579
25.0702 | -35.6609 | -35.264
25.0763 | -35.6614 | -35.2701
25.0824 | -35.6619 | -35.2759
25.0885 | -35.6622 | -35.2812
25.0946 | -35.6627 | -35.2862
25.1007 | -35.6631 | -35.2915
25.1068 | -35.6636 | -35.2971
25.1129 | -35.6639 | -35.303
25.119 | -35.6643 | -35.3083
25.1251 | -35.6648 | -35.314
25.1312 | -35.6652 | -35.3192
25.1373 | -35.6656 | -35.3249
25.1435 | -35.666 | -35.3306
25.1496 | -35.6663 | -35.3363
25.1557 | -35.6666 | -35.3419
25.1618 | -35.6671 | -35.3477
25.1679 | -35.6675 | -35.3535
25.174 | -35.668 | -35.359
25.1801 | -35.6684 | -35.3649
25.1862 | -35.6688 | -35.3708
25.1923 | -35.6692 | -35.3768
25.1984 | -35.6697 | -35.3825
25.2045 | -35.6701 | -35.3883
25.2106 | -35.6704 | -35.3943
25.2167 | -35.6707 | -35.4
25.2228 | -35.6711 | -35.406
25.2289 | -35.6716 | -35.4119
25.235 | -35.6719 | -35.4169
25.2411 | -35.6723 | -35.4227
25.2472 | -35.6728 | -35.4282
25.2533 | -35.6732 | -35.4343
25.2594 | -35.6735 | -35.4388
25.2655 | -35.674 | -35.4441
25.2716 | -35.6744 | -35.449
25.2777 | -35.6749 | -35.4542
25.2838 | -35.6753 | -35.4597
25.29 | -35.6757 | -35.4656
25.2961 | -35.6761 | -35.4719
25.3022 | -35.6766 | -35.4774
25.3083 | -35.677 | -35.4828
25.3144 | -35.6774 | -35.4879
25.3205 | -35.6777 | -35.4931
25.3266 | -35.6781 | -35.4985
25.3327 | -35.6786 | -35.5043
25.3388 | -35.679 | -35.51
25.3449 | -35.6795 | -35.5157
25.351 | -35.6799 | -35.5214
25.3571 | -35.6803 | -35.5279
25.3632 | -35.6807 | -35.5336
25.3693 | -35.6812 | -35.5393
25.3754 | -35.6815 | -35.5448
25.3815 | -35.682 | -35.5501
25.3876 | -35.6825 | -35.5557
25.3937 | -35.683 | -35.5613
25.3998 | -35.6834 | -35.5671
25.4059 | -35.6839 | -35.5731
25.412 | -35.6842 | -35.5788
25.4181 | -35.6847 | -35.5844
25.4242 | -35.6851 | -35.59
25.4304 | -35.6855 | -35.5953
25.4365 | -35.686 | -35.6013
25.4426 | -35.6864 | -35.6071
25.4487 | -35.6868 | -35.6133
25.4548 | -35.6872 | -35.6182
25.4609 | -35.6876 | -35.6235
25.467 | -35.6881 | -35.6305
25.4731 | -35.6884 | -35.6366
25.4792 | -35.6889 | -35.6422
25.4853 | -35.6893 | -35.6478
25.4914 | -35.6896 | -35.6533
25.4975 | -35.6901 | -35.6593
25.5036 | -35.6904 | -35.6652
25.5097 | -35.6909 | -35.6707
25.5158 | -35.6912 | -35.6769
25.5219 | -35.6917 | -35.6827
25.528 | -35.692 | -35.689
25.5341 | -35.6925 | -35.6941
25.5402 | -35.693 | -35.6996
25.5463 | -35.6933 | -35.7056
25.5524 | -35.6938 | -35.7118
25.5585 | -35.6942 | -35.7178
25.5646 | -35.6946 | -35.7236
25.5707 | -35.6949 | -35.7295
25.5769 | -35.6953 | -35.735
25.583 | -35.6957 | -35.7405
25.5891 | -35.6961 | -35.7463
25.5952 | -35.6965 | -35.7512
25.6013 | -35.697 | -35.7563
25.6074 | -35.6973 | -35.7611
25.6135 | -35.6977 | -35.7668
25.6196 | -35.6981 | -35.7724
25.6257 | -35.6984 | -35.7782
25.6318 | -35.6988 | -35.7834
25.6379 | -35.6993 | -35.7888
25.644 | -35.6997 | -35.7948
25.6501 | -35.7 | -35.8
25.6562 | -35.7004 | -35.8061
25.6623 | -35.7008 | -35.8111
25.6684 | -35.7013 | -35.8168
25.6745 | -35.7016 | -35.8225
25.6806 | -35.7021 | -35.8282
25.6867 | -35.7026 | -35.8337
25.6928 | -35.703 | -35.839
25.6989 | -35.7035 | -35.8449
25.705 | -35.7039 | -35.8507
25.7111 | -35.7042 | -35.8557
25.7173 | -35.7045 | -35.8609
25.7234 | -35.7049 | -35.8672
25.7295 | -35.7053 | -35.8728
25.7356 | -35.7056 | -35.879
25.7417 | -35.706 | -35.885
25.7478 | -35.7065 | -35.8906
25.7539 | -35.7069 | -35.8958
25.76 | -35.7074 | -35.9013
25.7661 | -35.7078 | -35.9068
25.7722 | -35.7082 | -35.9123
25.7783 | -35.7087 | -35.9177
25.7844 | -35.709 | -35.923
25.7905 | -35.7095 | -35.9287
25.7966 | -35.7099 | -35.9345
25.8027 | -35.7102 | -35.9396
25.8088 | -35.7106 | -35.9456
25.8149 | -35.711 | -35.9509
25.821 | -35.7114 | -35.9569
25.8271 | -35.7118 | -35.9625
25.8332 | -35.7122 | -35.9679
25.8393 | -35.7127 | -35.9738
25.8454 | -35.7131 | -35.979
25.8515 | -35.7135 | -35.9848
25.8576 | -35.7138 | -35.9913
25.8638 | -35.7142 | -35.9968
25.8699 | -35.7146 | -36.0021
25.876 | -35.715 | -36.0076
25.8821 | -35.7154 | -36.0136
25.8882 | -35.7156 | -36.0196
25.8943 | -35.716 | -36.0251
25.9004 | -35.7163 | -36.0309
25.9065 | -35.7167 | -36.0362
25.9126 | -35.7171 | -36.0418
25.9187 | -35.7176 | -36.0476
25.9248 | -35.718 | -36.0537
25.9309 | -35.7184 | -36.0589
25.937 | -35.7188 | -36.0649
25.9431 | -35.7191 | -36.0705
25.9492 | -35.7195 | -36.0759
25.9553 | -35.7199 | -36.0814
25.9614 | -35.7202 | -36.0871
25.9675 | -35.7206 | -36.0924
25.9736 | -35.721 | -36.0982
25.9797 | -35.7214 | -36.1042
25.9858 | -35.7218 | -36.11
25.9919 | -35.7221 | -36.1161
25.998 | -35.7225 | -36.1209
26.0042 | -35.7228 | -36.1268
26.0103 | -35.7232 | -36.1323
26.0164 | -35.7235 | -36.1381
26.0225 | -35.724 | -36.1436
26.0286 | -35.7244 | -36.1491
26.0347 | -35.7247 | -36.1555
26.0408 | -35.7251 | -36.1609
26.0469 | -35.7255 | -36.1667
26.053 | -35.7259 | -36.1722
26.0591 | -35.7264 | -36.1773
26.0652 | -35.7266 | -36.1831
26.0713 | -35.727 | -36.1881
26.0774 | -35.7275 | -36.1941
26.0835 | -35.7278 | -36.1999
26.0896 | -35.7282 | -36.2049
26.0957 | -35.7286 | -36.2108
26.1018 | -35.729 | -36.2165
26.1079 | -35.7292 | -36.2217
26.114 | -35.7296 | -36.228
26.1201 | -35.73 | -36.2336
26.1262 | -35.7303 | -36.2391
26.1323 | -35.7307 | -36.2444
26.1384 | -35.731 | -36.2504
26.1445 | -35.7314 | -36.2561
26.1507 | -35.7318 | -36.2615
26.1568 | -35.7321 | -36.2671
26.1629 | -35.7326 | -36.273
26.169 | -35.733 | -36.2788
26.1751 | -35.7334 | -36.2849
26.1812 | -35.7338 | -36.2905
26.1873 | -35.7341 | -36.2963
26.1934 | -35.7345 | -36.3022
26.1995 | -35.735 | -36.3077
26.2056 | -35.7353 | -36.3133
26.2117 | -35.7357 | -36.3192
26.2178 | -35.7361 | -36.3247
26.2239 | -35.7365 | -36.3305
26.23 | -35.7368 | -36.3363
26.2361 | -35.7373 | -36.3417
26.2422 | -35.7376 | -36.3473
26.2483 | -35.7379 | -36.3529
26.2544 | -35.7383 | -36.3587
26.2605 | -35.7387 | -36.3643
26.2666 | -35.7391 | -36.37
26.2727 | -35.7395 | -36.3756
26.2788 | -35.7398 | -36.3818
26.2849 | -35.7402 | -36.387
26.2911 | -35.7406 | -36.3928
26.2972 | -35.741 | -36.3988
26.3033 | -35.7413 | -36.4046
26.3094 | -35.7417 | -36.4104
26.3155 | -35.7421 | -36.4159
26.3216 | -35.7424 | -36.4216
26.3277 | -35.7428 | -36.427
26.3338 | -35.7431 | -36.4323
26.3399 | -35.7435 | -36.438
26.346 | -35.7439 | -36.4438
26.3521 | -35.7443 | -36.4492
26.3582 | -35.7445 | -36.4548
26.3643 | -35.7449 | -36.4604
26.3704 | -35.7453 | -36.4663
26.3765 | -35.7458 | -36.4718
26.3826 | -35.7461 | -36.4776
26.3887 | -35.7465 | -36.4835
26.3948 | -35.7468 | -36.4891
26.4009 | -35.7471 | -36.4946
26.407 | -35.7475 | -36.5005
26.4131 | -35.7478 | -36.5061
26.4192 | -35.7482 | -36.5122
26.4253 | -35.7486 | -36.5183
26.4314 | -35.749 | -36.5237
26.4376 | -35.7495 | -36.529
26.4437 | -35.7499 | -36.5343
26.4498 | -35.7502 | -36.5403
26.4559 | -35.7506 | -36.5462
26.462 | -35.751 | -36.5513
26.4681 | -35.7514 | -36.5568
26.4742 | -35.7517 | -36.563
26.4803 | -35.752 | -36.5687
26.4864 | -35.7523 | -36.5739
26.4925 | -35.7527 | -36.5791
26.4986 | -35.753 | -36.5848
26.5047 | -35.7534 | -36.5907
26.5108 | -35.7539 | -36.5963
26.5169 | -35.7542 | -36.6021
26.523 | -35.7545 | -36.6075
26.5291 | -35.7549 | -36.6126
26.5352 | -35.7552 | -36.6181
26.5413 | -35.7556 | -36.624
26.5474 | -35.756 | -36.6302
26.5535 | -35.7564 | -36.636
26.5596 | -35.7567 | -36.6418
26.5657 | -35.7571 | -36.6473
26.5718 | -35.7574 | -36.6531
26.578 | -35.7578 | -36.6588
26.5841 | -35.7581 | -36.6648
26.5902 | -35.7585 | -36.6701
26.5963 | -35.7589 | -36.6758
26.6024 | -35.7592 | -36.6814
26.6085 | -35.7596 | -36.6873
26.6146 | -35.76 | -36.6928
26.6207 | -35.7604 | -36.6988
26.6268 | -35.7608 | -36.7042
26.6329 | -35.7612 | -36.7102
26.639 | -35.7615 | -36.7155
26.6451 | -35.7618 | -36.721
26.6512 | -35.7622 | -36.7263
26.6573 | -35.7626 | -36.7318
26.6634 | -35.7628 | -36.7376
26.6695 | -35.7632 | -36.7436
26.6756 | -35.7636 | -36.7492
26.6817 | -35.7639 | -36.7554
26.6878 | -35.7642 | -36.7611
26.6939 | -35.7646 | -36.7668
26.7 | -35.765 | -36.7729
26.7061 | -35.7654 | -36.7786
26.7122 | -35.7657 | -36.7847
26.7183 | -35.7661 | -36.7905
26.7245 | -35.7665 | -36.7962
26.7306 | -35.7669 | -36.8016
26.7367 | -35.7672 | -36.8073
26.7428 | -35.7676 | -36.8127
26.7489 | -35.768 | -36.8183
26.755 | -35.7684 | -36.8241
26.7611 | -35.7688 | -36.8298
26.7672 | -35.7692 | -36.8355
26.7733 | -35.7695 | -36.8412
26.7794 | -35.7699 | -36.847
26.7855 | -35.7702 | -36.8528
26.7916 | -35.7706 | -36.8584
26.7977 | -35.771 | -36.8639
26.8038 | -35.7714 | -36.8694
26.8099 | -35.7718 | -36.8753
26.816 | -35.7721 | -36.8809
26.8221 | -35.7724 | -36.8865
26.8282 | -35.7728 | -36.8919
26.8343 | -35.7732 | -36.8976
26.8404 | -35.7735 | -36.9038
26.8465 | -35.7739 | -36.9095
26.8526 | -35.7743 | -36.9152
26.8587 | -35.7747 | -36.9209
26.8649 | -35.7749 | -36.9266
26.871 | -35.7753 | -36.9323
26.8771 | -35.7758 | -36.9384
26.8832 | -35.7761 | -36.9443
26.8893 | -35.7765 | -36.9497
26.8954 | -35.7769 | -36.9557
26.9015 | -35.7772 | -36.9616
26.9076 | -35.7775 | -36.9673
26.9137 | -35.7778 | -36.9726
26.9198 | -35.7782 | -36.9786
26.9259 | -35.7785 | -36.984
26.932 | -35.7788 | -36.9897
26.9381 | -35.7793 | -36.9956
26.9442 | -35.7796 | -37.0012
26.9503 | -35.7799 | -37.0065
26.9564 | -35.7803 | -37.0122
26.9625 | -35.7807 | -37.0179
26.9686 | -35.7811 | -37.0234
26.9747 | -35.7815 | -37.0292
26.9808 | -35.7819 | -37.0349
26.9869 | -35.7823 | -37.0401
26.993 | -35.7826 | -37.0453
26.9991 | -35.783 | -37.0505
27.0052 | -35.7834 | -37.0562
27.0114 | -35.7838 | -37.0623
27.0175 | -35.7842 | -37.0682
27.0236 | -35.7846 | -37.0739
27.0297 | -35.7849 | -37.0798
27.0358 | -35.7853 | -37.0855
27.0419 | -35.7857 | -37.0916
27.048 | -35.786 | -37.0974
27.0541 | -35.7864 | -37.1029
27.0602 | -35.7868 | -37.1088
27.0663 | -35.7872 | -37.1143
27.0724 | -35.7875 | -37.1202
27.0785 | -35.7879 | -37.1258
27.0846 | -35.7882 | -37.1315
27.0907 | -35.7886 | -37.1375
27.0968 | -35.7889 | -37.1432
27.1029 | -35.7893 | -37.1489
27.109 | -35.7896 | -37.1543
27.1151 | -35.79 | -37.1599
27.1212 | -35.7903 | -37.1654
27.1273 | -35.7907 | -37.1712
27.1334 | -35.791 | -37.1773
27.1395 | -35.7914 | -37.1833
27.1456 | -35.7918 | -37.189
27.1518 | -35.7921 | -37.1947
27.1579 | -35.7924 | -37.1997
27.164 | -35.7927 | -37.2051
27.1701 | -35.7931 | -37.2104
27.1762 | -35.7935 | -37.2156
27.1823 | -35.7938 | -37.2211
27.1884 | -35.7942 | -37.2272
27.1945 | -35.7946 | -37.2323
27.2006 | -35.7949 | -37.2379
27.2067 | -35.7953 | -37.2434
27.2128 | -35.7956 | -37.2488
27.2189 | -35.796 | -37.2547
27.225 | -35.7964 | -37.2604
27.2311 | -35.7967 | -37.2655
27.2372 | -35.7971 | -37.2711
27.2433 | -35.7974 | -37.2764
27.2494 | -35.7978 | -37.2824
27.2555 | -35.7982 | -37.2876
27.2616 | -35.7985 | -37.2938
27.2677 | -35.7989 | -37.2994
27.2738 | -35.7993 | -37.3055
27.2799 | -35.7996 | -37.3114
27.286 | -35.8 | -37.3169
27.2921 | -35.8004 | -37.323
27.2983 | -35.8008 | -37.328
27.3044 | -35.8011 | -37.3333
27.3105 | -35.8015 | -37.3391
27.3166 | -35.8018 | -37.3448
27.3227 | -35.8022 | -37.3506
27.3288 | -35.8025 | -37.3563
27.3349 | -35.8029 | -37.3619
27.341 | -35.8032 | -37.3676
27.3471 | -35.8036 | -37.3728
27.3532 | -35.8039 | -37.378
27.3593 | -35.8043 | -37.3837
27.3654 | -35.8047 | -37.3894
27.3715 | -35.805 | -37.3946
27.3776 | -35.8054 | -37.4005
27.3837 | -35.8058 | -37.4059
27.3898 | -35.8062 | -37.4113
27.3959 | -35.8065 | -37.4166
27.402 | -35.8069 | -37.4225
27.4081 | -35.8073 | -37.4282
27.4142 | -35.8076 | -37.4338
27.4203 | -35.808 | -37.4392
27.4264 | -35.8083 | -37.445
27.4325 | -35.8087 | -37.4508
27.4387 | -35.809 | -37.4568
27.4448 | -35.8093 | -37.4624
27.4509 | -35.8096 | -37.4681
27.457 | -35.81 | -37.4737
27.4631 | -35.8104 | -37.4794
27.4692 | -35.8106 | -37.4848
27.4753 | -35.811 | -37.4902
27.4814 | -35.8113 | -37.4958
27.4875 | -35.8116 | -37.5017
27.4936 | -35.812 | -37.5072
27.4997 | -35.8124 | -37.5127
27.5058 | -35.8127 | -37.5185
27.5119 | -35.813 | -37.5242
27.518 | -35.8134 | -37.5299
27.5241 | -35.8138 | -37.5356
27.5302 | -35.8141 | -37.541
27.5363 | -35.8144 | -37.5466
27.5424 | -35.8147 | -37.552
27.5485 | -35.8151 | -37.5581
27.5546 | -35.8155 | -37.5636
27.5607 | -35.8158 | -37.5689
27.5668 | -35.8161 | -37.5743
27.5729 | -35.8165 | -37.5798
27.5791 | -35.8168 | -37.585
27.5852 | -35.8171 | -37.5906
27.5913 | -35.8175 | -37.5962
27.5974 | -35.8179 | -37.6023
27.6035 | -35.8183 | -37.6078
27.6096 | -35.8185 | -37.6136
27.6157 | -35.819 | -37.6195
27.6218 | -35.8193 | -37.625
27.6279 | -35.8197 | -37.6303
27.634 | -35.8201 | -37.6361
27.6401 | -35.8204 | -37.6412
27.6462 | -35.8208 | -37.6468
27.6523 | -35.8211 | -37.6523
27.6584 | -35.8214 | -37.6578
27.6645 | -35.8218 | -37.6634
27.6706 | -35.8221 | -37.6693
27.6767 | -35.8225 | -37.6749
27.6828 | -35.8228 | -37.6803
27.6889 | -35.823 | -37.6854
27.695 | -35.8234 | -37.691
27.7011 | -35.8238 | -37.6964
27.7072 | -35.8242 | -37.7021
27.7133 | -35.8245 | -37.7077
27.7194 | -35.8249 | -37.7132
27.7256 | -35.8252 | -37.7183
27.7317 | -35.8255 | -37.7235
27.7378 | -35.8259 | -37.729
27.7439 | -35.8262 | -37.7349
27.75 | -35.8266 | -37.7408
27.7561 | -35.8269 | -37.7458
27.7622 | -35.8272 | -37.7514
27.7683 | -35.8276 | -37.7573
27.7744 | -35.8279 | -37.7627
27.7805 | -35.8283 | -37.7681
27.7866 | -35.8287 | -37.7737
27.7927 | -35.829 | -37.779
27.7988 | -35.8293 | -37.7844
27.8049 | -35.8297 | -37.7898
27.811 | -35.83 | -37.7953
27.8171 | -35.8303 | -37.8009
27.8232 | -35.8307 | -37.8062
27.8293 | -35.831 | -37.8117
27.8354 | -35.8314 | -37.8176
27.8415 | -35.8317 | -37.8229
27.8476 | -35.8321 | -37.8287
27.8537 | -35.8324 | -37.8344
27.8598 | -35.8327 | -37.8394
27.866 | -35.8331 | -37.8449
27.8721 | -35.8335 | -37.8502
27.8782 | -35.8338 | -37.8558
27.8843 | -35.8342 | -37.8615
27.8904 | -35.8345 | -37.867
27.8965 | -35.835 | -37.8724
27.9026 | -35.8353 | -37.8779
27.9087 | -35.8357 | -37.8834
27.9148 | -35.836 | -37.8885
27.9209 | -35.8364 | -37.8942
27.927 | -35.8368 | -37.8992
27.9331 | -35.8372 | -37.9044
27.9392 | -35.8375 | -37.9098
27.9453 | -35.8379 | -37.9153
27.9514 | -35.8383 | -37.9208
27.9575 | -35.8386 | -37.9263
27.9636 | -35.839 | -37.9315
27.9697 | -35.8394 | -37.9371
27.9758 | -35.8397 | -37.9425
27.9819 | -35.84 | -37.9483
27.988 | -35.8404 | -37.9538
27.9941 | -35.8407 | -37.9594
28.0002 | -35.8411 | -37.9646
28.0063 | -35.8415 | -37.9699
28.0125 | -35.8418 | -37.9754
28.0186 | -35.8422 | -37.9809
28.0247 | -35.8425 | -37.9862
28.0308 | -35.8429 | -37.9914
28.0369 | -35.8432 | -37.9967
28.043 | -35.8436 | -38.0025
28.0491 | -35.844 | -38.0082
28.0552 | -35.8443 | -38.0136
28.0613 | -35.8447 | -38.0191
28.0674 | -35.8451 | -38.0245
28.0735 | -35.8454 | -38.0299

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4332 | -41.9059
21.4908 | -35.5863 | -34.771
16.4948 | -34.9338 | -27.4831
12.6602 | -34.5643 | -21.2735
9.7171 | -34.3488 | -16.3984
7.4582 | -34.195 | -12.4468
5.7244 | -34.0986 | -9.24
4.3936 | -34.0311 | -6.6521
3.3722 | -34.0023 | -4.6123
2.5883 | -33.9924 | -3.0666
1.9866 | -33.9698 | -1.8325
1.5248 | -33.975 | -0.84496
1.1703 | -33.9827 | -0.19991
1 | -33.9923 | 0.1184
0.89822 | -34.0026 | 0.38944
0.68941 | -34.0141 | 0.96631
0.52919 | -34.0363 | 1.4134
0.40612 | -34.0394 | 2.0148
0.31171 | -34.0227 | 2.4625
0.23928 | -34.0078 | 3.1794
0.18358 | -34.011 | 3.5923
0.14091 | -34.031 | 4.4681
0.10819 | -34.0474 | 5.6287
0.083017 | -34.0468 | 7.1098
0.063705 | -34.0983 | 8.9655
0.048923 | -34.1606 | 11.4374
0.037479 | -34.2663 | 14.6367
0.028801 | -34.4104 | 18.6778
0.022078 | -34.7088 | 24.111
0.017023 | -35.1963 | 30.5563
0.013018 | -36.0049 | 38.2512
0.010014 | -37.1286 | 46.1801

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003688 | -17.4141
0.007377 | -33.4082
0.011065 | -37.3994
0.014753 | -40.3359
0.018442 | -43.3032
0.02213 | -45.8141
0.025818 | -47.5122
0.029507 | -48.5013
0.033195 | -49.398
0.036883 | -51.1543
0.040572 | -52.2384
0.04426 | -53.2629
0.047948 | -53.411
0.051637 | -55.6576
0.055325 | -56.573
0.059013 | -55.6435
0.062702 | -56.2255
0.06639 | -58.1932
0.070078 | -58.6332
0.073767 | -59.5317
0.077455 | -60.118
0.081143 | -61.5113
0.084832 | -61.9421
0.08852 | -62.0394
0.092208 | -61.8742
0.095897 | -61.2463
0.099585 | -62.6414
0.10327 | -63.3908
0.10696 | -63.4013
0.11065 | -64.1724
0.11434 | -64.6247
0.11803 | -64.3941
0.12172 | -64.3412
0.1254 | -65.5646
0.12909 | -67.0489
0.13278 | -66.6419
0.13647 | -65.8636
0.14016 | -66.0016
0.14385 | -66.0713
0.14753 | -65.9144
0.15122 | -66.7688
0.15491 | -67.8199
0.1586 | -68.2062
0.16229 | -68.1566
0.16598 | -69.0944
0.16966 | -69.5682
0.17335 | -69.6495
0.17704 | -69.0381
0.18073 | -69.9929
0.18442 | -70.6178
0.1881 | -71.3595
0.19179 | -72.7113
0.19548 | -71.3705
0.19917 | -70.1517
0.20286 | -71.0112
0.20655 | -71.9228
0.21024 | -71.6661
0.21392 | -72.4285
0.21761 | -72.6921
0.2213 | -72.5934
0.22499 | -72.1368
0.22868 | -72.7399
0.23236 | -73.6153
0.23605 | -73.2237
0.23974 | -72.212
0.24343 | -73.6187
0.24712 | -74.4388
0.25081 | -74.3497
0.2545 | -74.4898
0.25818 | -74.7832
0.26187 | -75.78
0.26556 | -74.7583
0.26925 | -73.9108
0.27294 | -73.945
0.27663 | -75.1584
0.28031 | -74.952
0.284 | -74.4469
0.28769 | -74.9305
0.29138 | -75.0644
0.29507 | -75.1986
0.29875 | -76.0535
0.30244 | -76.4044
0.30613 | -76.1801
0.30982 | -75.2051
0.31351 | -75.5882
0.3172 | -76.556
0.32088 | -75.4167
0.32457 | -75.1091
0.32826 | -75.8317
0.33195 | -77.1308
0.33564 | -78.2729
0.33933 | -75.8986
0.34302 | -76.9768
0.3467 | -77.8937
0.35039 | -77.4785
0.35408 | -78.0373
0.35777 | -78.175
0.36146 | -78.2049
0.36514 | -78.7302
0.36883 | -80.0309
0.37252 | -79.1113
0.37621 | -77.4373
0.3799 | -77.3424
0.38359 | -77.3876
0.38727 | -77.3734
0.39096 | -77.8442
0.39465 | -78.2416
0.39834 | -78.1
0.40203 | -77.8059
0.40572 | -78.903
0.40941 | -78.2765
0.41309 | -77.5362
0.41678 | -77.8783
0.42047 | -78.3958
0.42416 | -78.6677
0.42785 | -78.6853
0.43154 | -78.7455
0.43522 | -79.0614
0.43891 | -78.4826
0.4426 | -79.1431
0.44629 | -79.3497
0.44998 | -79.4014
0.45366 | -80.1013
0.45735 | -79.5374
0.46104 | -80.3243
0.46473 | -81.2965
0.46842 | -80.2739
0.47211 | -79.0698
0.4758 | -79.3726
0.47948 | -79.9974
0.48317 | -80.264
0.48686 | -80.5794
0.49055 | -80.6292
0.49424 | -81.3646
0.49793 | -81.1416
0.50161 | -79.6845
0.5053 | -79.8797
0.50899 | -79.9591
0.51268 | -80.1297
0.51637 | -80.6083
0.52006 | -80.4892
0.52374 | -80.2546
0.52743 | -79.5376
0.53112 | -80.4017
0.53481 | -81.0219
0.5385 | -81.5291
0.54219 | -81.9611
0.54587 | -82.2105
0.54956 | -82.2296
0.55325 | -82.0058
0.55694 | -81.6424
0.56063 | -81.4606
0.56432 | -81.6409
0.568 | -81.6844
0.57169 | -81.5902
0.57538 | -81.6166
0.57907 | -81.5878
0.58276 | -81.6378
0.58644 | -81.96
0.59013 | -82.2008
0.59382 | -82.5357
0.59751 | -82.8097
0.6012 | -82.7914
0.60489 | -82.8144
0.60857 | -83.0319
0.61226 | -83.0114
0.61595 | -82.4777
0.61964 | -82.2552
0.62333 | -82.4501
0.62702 | -82.6027
0.6307 | -82.9036
0.63439 | -83.574
0.63808 | -83.8433
0.64177 | -83.3446
0.64546 | -82.957
0.64915 | -82.8757
0.65284 | -83.159
0.65652 | -83.6656
0.66021 | -83.6162
0.6639 | -82.961
0.66759 | -82.7715
0.67128 | -83.5106
0.67497 | -84.1832
0.67865 | -84.5493
0.68234 | -84.6223
0.68603 | -84.3331
0.68972 | -84.009
0.69341 | -83.689
0.6971 | -83.5208
0.70078 | -83.7342
0.70447 | -84.066
0.70816 | -84.0496
0.71185 | -84.0772
0.71554 | -84.0436
0.71923 | -83.8065
0.72291 | -83.7295
0.7266 | -83.7472
0.73029 | -83.6829
0.73398 | -83.4993
0.73767 | -83.8951
0.74135 | -84.914
0.74504 | -85.5166
0.74873 | -85.1519
0.75242 | -84.5155
0.75611 | -84.114
0.7598 | -83.9938
0.76348 | -83.9785
0.76717 | -84.4898
0.77086 | -84.8995
0.77455 | -85.1389
0.77824 | -85.4773
0.78193 | -85.4002
0.78561 | -84.7084
0.7893 | -84.0175
0.79299 | -84.1829
0.79668 | -85.1089
0.80037 | -85.8059
0.80406 | -86.054
0.80775 | -86.0071
0.81143 | -85.811
0.81512 | -86.081
0.81881 | -86.3321
0.8225 | -85.7456
0.82619 | -85.0844
0.82988 | -85.2431
0.83356 | -85.8399
0.83725 | -85.8721
0.84094 | -85.2764
0.84463 | -85.1044
0.84832 | -85.2271
0.85201 | -85.2098
0.85569 | -85.3645
0.85938 | -85.6065
0.86307 | -85.4664
0.86676 | -85.2487
0.87045 | -85.3904
0.87414 | -85.5938
0.87782 | -86.0532
0.88151 | -86.4675
0.8852 | -86.3817
0.88889 | -86.2259
0.89258 | -86.1312
0.89627 | -86.119
0.89995 | -86.1623
0.90364 | -85.9981
0.90733 | -86.0403
0.91102 | -86.2541
0.91471 | -86.3194
0.9184 | -86.277
0.92208 | -86.3708
0.92577 | -86.2733
0.92946 | -86.0562
0.93315 | -85.9676
0.93684 | -86.1657
0.94053 | -86.4661
0.94421 | -86.8166
0.9479 | -87.157
0.95159 | -87.1646
0.95528 | -86.7793
0.95897 | -86.0789
0.96266 | -85.6452
0.96634 | -86.1215
0.97003 | -86.7612
0.97372 | -86.8064
0.97741 | -87.0129
0.9811 | -87.2814
0.98479 | -86.919
0.98847 | -86.462
0.99216 | -86.3603
0.99585 | -86.4871
0.99954 | -86.9203
1.0032 | -87.4864
1.0069 | -87.5036
1.0106 | -87.4126
1.0143 | -87.3754
1.018 | -87.4498
1.0217 | -87.5646
1.0254 | -87.5524
1.029 | -87.5145
1.0327 | -87.8829
1.0364 | -88.1171
1.0401 | -88.0502
1.0438 | -88.2774
1.0475 | -88.1683
1.0512 | -87.8708
1.0549 | -88.0988
1.0586 | -88.33
1.0622 | -87.939
1.0659 | -87.6572
1.0696 | -87.5184
1.0733 | -87.5625
1.077 | -87.72
1.0807 | -87.9355
1.0844 | -88.2323
1.0881 | -88.5973
1.0917 | -88.7621
1.0954 | -88.6021
1.0991 | -88.4098
1.1028 | -88.2851
1.1065 | -88.1626
1.1102 | -88.099
1.1139 | -88.1428
1.1176 | -88.2394
1.1213 | -88.2961
1.1249 | -88.286
1.1286 | -88.1976
1.1323 | -88.1137
1.136 | -88.2905
1.1397 | -88.77
1.1434 | -89.2479
1.1471 | -89.3821
1.1508 | -89.1837
1.1544 | -89.0058
1.1581 | -89.0004
1.1618 | -89.0517
1.1655 | -89.0289
1.1692 | -88.8808
1.1729 | -88.6971
1.1766 | -88.6055
1.1803 | -88.637
1.184 | -88.778
1.1876 | -88.9813
1.1913 | -89.1738
1.195 | -89.2631
1.1987 | -89.1819
1.2024 | -88.9574
1.2061 | -88.7086
1.2098 | -88.5334
1.2135 | -88.3727
1.2172 | -88.2093
1.2208 | -88.3175
1.2245 | -88.791
1.2282 | -89.1792
1.2319 | -89.1484
1.2356 | -89.0804
1.2393 | -89.3075
1.243 | -89.6801
1.2467 | -89.8027
1.2503 | -89.742
1.254 | -89.8445
1.2577 | -90.0645
1.2614 | -90.114
1.2651 | -90.0022
1.2688 | -89.9568
1.2725 | -89.9316
1.2762 | -89.6302
1.2799 | -89.1074
1.2835 | -88.6875
1.2872 | -88.5759
1.2909 | -88.7238
1.2946 | -88.9898
1.2983 | -89.2649
1.302 | -89.3735
1.3057 | -89.275
1.3094 | -89.178
1.313 | -89.258
1.3167 | -89.5263
1.3204 | -89.8029
1.3241 | -89.9181
1.3278 | -89.9779
1.3315 | -90.0911
1.3352 | -90.1637
1.3389 | -90.1024
1.3426 | -89.943
1.3462 | -89.8854
1.3499 | -90.1769
1.3536 | -90.6773
1.3573 | -90.7717
1.361 | -90.3412
1.3647 | -89.964
1.3684 | -89.9213
1.3721 | -90.0738
1.3757 | -90.2781
1.3794 | -90.4984
1.3831 | -90.5809
1.3868 | -90.3401
1.3905 | -90.0289
1.3942 | -90.0337
1.3979 | -90.2698
1.4016 | -90.3477
1.4053 | -90.2671
1.4089 | -90.2655
1.4126 | -90.2955
1.4163 | -90.2002
1.42 | -90.0238
1.4237 | -89.8176
1.4274 | -89.6449
1.4311 | -89.6419
1.4348 | -89.8655
1.4385 | -90.2671
1.4421 | -90.6866
1.4458 | -90.7775
1.4495 | -90.3488
1.4532 | -89.7161
1.4569 | -89.354
1.4606 | -89.4905
1.4643 | -90.076
1.468 | -90.8209
1.4716 | -91.2756
1.4753 | -91.2743
1.479 | -90.9787
1.4827 | -90.508
1.4864 | -90.0577
1.4901 | -89.8609
1.4938 | -89.9467
1.4975 | -90.2094
1.5012 | -90.5805
1.5048 | -91.0135
1.5085 | -91.3393
1.5122 | -91.4422
1.5159 | -91.4748
1.5196 | -91.6194
1.5233 | -91.8709
1.527 | -92.1144
1.5307 | -92.25
1.5343 | -92.2719
1.538 | -92.2231
1.5417 | -92.1171
1.5454 | -91.9275
1.5491 | -91.6544
1.5528 | -91.3662
1.5565 | -91.1677
1.5602 | -91.1135
1.5639 | -91.1604
1.5675 | -91.267
1.5712 | -91.4212
1.5749 | -91.4668
1.5786 | -91.341
1.5823 | -91.2983
1.586 | -91.4883
1.5897 | -91.7415
1.5934 | -91.7714
1.597 | -91.8374
1.6007 | -91.8921
1.6044 | -91.8931
1.6081 | -91.757
1.6118 | -91.5256
1.6155 | -91.3092
1.6192 | -91.1422
1.6229 | -91.008
1.6266 | -90.9095
1.6302 | -90.8878
1.6339 | -90.9499
1.6376 | -91.0331
1.6413 | -91.0936
1.645 | -91.1429
1.6487 | -91.2398
1.6524 | -91.3773
1.6561 | -91.4884
1.6598 | -91.5161
1.6634 | -91.4501
1.6671 | -91.3918
1.6708 | -91.4695
1.6745 | -91.7147
1.6782 | -92.0058
1.6819 | -92.1561
1.6856 | -92.101
1.6893 | -91.9655
1.6929 | -91.9253
1.6966 | -92.0457
1.7003 | -92.2327
1.704 | -92.3845
1.7077 | -92.4512
1.7114 | -92.4319
1.7151 | -92.3543
1.7188 | -92.2863
1.7225 | -92.2843
1.7261 | -92.3427
1.7298 | -92.4008
1.7335 | -92.4009
1.7372 | -92.3451
1.7409 | -92.2053
1.7446 | -91.9162
1.7483 | -91.5226
1.752 | -91.1704
1.7556 | -90.9639
1.7593 | -90.95
1.763 | -91.13
1.7667 | -91.4005
1.7704 | -91.5862
1.7741 | -91.6156
1.7778 | -91.5549
1.7815 | -91.5158
1.7852 | -91.562
1.7888 | -91.6926
1.7925 | -91.8211
1.7962 | -91.8936
1.7999 | -91.979
1.8036 | -92.1703
1.8073 | -92.4791
1.811 | -92.7847
1.8147 | -92.968
1.8183 | -93.0226
1.822 | -92.988
1.8257 | -92.9192
1.8294 | -92.8268
1.8331 | -92.7226
1.8368 | -92.6265
1.8405 | -92.5839
1.8442 | -92.6704
1.8479 | -92.891
1.8515 | -93.1363
1.8552 | -93.1956
1.8589 | -92.9118
1.8626 | -92.4121
1.8663 | -92.009
1.87 | -91.9339
1.8737 | -92.1832
1.8774 | -92.5606
1.8811 | -92.8574
1.8847 | -92.9951
1.8884 | -93.0236
1.8921 | -93.0235
1.8958 | -93.0167
1.8995 | -92.9625
1.9032 | -92.8173
1.9069 | -92.597
1.9106 | -92.3703
1.9142 | -92.2339
1.9179 | -92.2506
1.9216 | -92.3789
1.9253 | -92.5097
1.929 | -92.6294
1.9327 | -92.8083
1.9364 | -93.0587
1.9401 | -93.2874
1.9438 | -93.3635
1.9474 | -93.2609
1.9511 | -93.1134
1.9548 | -93.0453
1.9585 | -93.0848
1.9622 | -93.1795
1.9659 | -93.327
1.9696 | -93.4971
1.9733 | -93.5821
1.9769 | -93.4982
1.9806 | -93.2922
1.9843 | -93.0869
1.988 | -92.966
1.9917 | -92.926
1.9954 | -92.9349
1.9991 | -92.9784
2.0028 | -93.0608
2.0065 | -93.1922
2.0101 | -93.3572
2.0138 | -93.5516
2.0175 | -93.7873
2.0212 | -94.0267
2.0249 | -94.1529
2.0286 | -94.0544
2.0323 | -93.7983
2.036 | -93.5309
2.0396 | -93.3172
2.0433 | -93.1773
2.047 | -93.0486
2.0507 | -92.9185
2.0544 | -92.8772
2.0581 | -93.0195
2.0618 | -93.3221
2.0655 | -93.5941
2.0692 | -93.6826
2.0728 | -93.6221
2.0765 | -93.5456
2.0802 | -93.5529
2.0839 | -93.6127
2.0876 | -93.6797
2.0913 | -93.7597
2.095 | -93.8827
2.0987 | -94.022
2.1024 | -94.0894
2.106 | -94.04
2.1097 | -93.8878
2.1134 | -93.7154
2.1171 | -93.6126
2.1208 | -93.5781
2.1245 | -93.5724
2.1282 | -93.5633
2.1319 | -93.5627
2.1355 | -93.5746
2.1392 | -93.6074
2.1429 | -93.6597
2.1466 | -93.7184
2.1503 | -93.725
2.154 | -93.667
2.1577 | -93.5915
2.1614 | -93.5366
2.1651 | -93.5165
2.1687 | -93.5214
2.1724 | -93.5171
2.1761 | -93.5116
2.1798 | -93.5328
2.1835 | -93.6143
2.1872 | -93.7587
2.1909 | -93.9282
2.1946 | -94.0712
2.1982 | -94.1585
2.2019 | -94.1936
2.2056 | -94.1919
2.2093 | -94.192
2.213 | -94.2435
2.2167 | -94.3525
2.2204 | -94.47
2.2241 | -94.5363
2.2278 | -94.5338
2.2314 | -94.4688
2.2351 | -94.3492
2.2388 | -94.195
2.2425 | -94.0287
2.2462 | -93.9044
2.2499 | -93.8371
2.2536 | -93.8124
2.2573 | -93.8014
2.2609 | -93.7866
2.2646 | -93.771
2.2683 | -93.761
2.272 | -93.7828
2.2757 | -93.8344
2.2794 | -93.9061
2.2831 | -93.9881
2.2868 | -94.0854
2.2905 | -94.2473
2.2941 | -94.4836
2.2978 | -94.7405
2.3015 | -94.9359
2.3052 | -94.9957
2.3089 | -94.9053
2.3126 | -94.7095
2.3163 | -94.4955
2.32 | -94.3272
2.3237 | -94.2635
2.3273 | -94.3187
2.331 | -94.4696
2.3347 | -94.6482
2.3384 | -94.7595
2.3421 | -94.7165
2.3458 | -94.5181
2.3495 | -94.2437
2.3532 | -93.9492
2.3568 | -93.6913
2.3605 | -93.5511
2.3642 | -93.5945
2.3679 | -93.8342
2.3716 | -94.2012
2.3753 | -94.5781
2.379 | -94.8764
2.3827 | -95.073
2.3864 | -95.1636
2.39 | -95.1391
2.3937 | -95.024
2.3974 | -94.8757
2.4011 | -94.7708
2.4048 | -94.743
2.4085 | -94.7946
2.4122 | -94.8683
2.4159 | -94.893
2.4195 | -94.8148
2.4232 | -94.6543
2.4269 | -94.47
2.4306 | -94.3411
2.4343 | -94.3213
2.438 | -94.41
2.4417 | -94.5805
2.4454 | -94.7749
2.4491 | -94.91
2.4527 | -94.9639
2.4564 | -94.9913
2.4601 | -95.0317
2.4638 | -95.1015
2.4675 | -95.1831
2.4712 | -95.2475
2.4749 | -95.2895
2.4786 | -95.2781
2.4823 | -95.1883
2.4859 | -95.076
2.4896 | -94.9829
2.4933 | -94.9527
2.497 | -94.9989
2.5007 | -95.1215
2.5044 | -95.2685
2.5081 | -95.361
2.5118 | -95.3487
2.5154 | -95.2491
2.5191 | -95.1454
2.5228 | -95.0844
2.5265 | -95.0776
2.5302 | -95.13
2.5339 | -95.2133
2.5376 | -95.2889
2.5413 | -95.3303
2.545 | -95.3214
2.5486 | -95.2806
2.5523 | -95.2224
2.556 | -95.1652
2.5597 | -95.1214
2.5634 | -95.1026
2.5671 | -95.1075
2.5708 | -95.1532
2.5745 | -95.2396
2.5781 | -95.351
2.5818 | -95.4533
2.5855 | -95.5302
2.5892 | -95.5712
2.5929 | -95.5727
2.5966 | -95.5285
2.6003 | -95.4592
2.604 | -95.4121
2.6077 | -95.4431
2.6113 | -95.5712
2.615 | -95.7667
2.6187 | -95.973
2.6224 | -96.1267
2.6261 | -96.1919
2.6298 | -96.1596
2.6335 | -96.0493
2.6372 | -95.9293
2.6408 | -95.8764
2.6445 | -95.9181
2.6482 | -95.99
2.6519 | -96.0182
2.6556 | -95.9846
2.6593 | -95.8631
2.663 | -95.7762
2.6667 | -95.6966
2.6704 | -95.6428
2.674 | -95.6223
2.6777 | -95.6319
2.6814 | -95.6542
2.6851 | -95.6683
2.6888 | -95.6902
2.6925 | -95.755
2.6962 | -95.8525
2.6999 | -95.9474
2.7035 | -96.0041
2.7072 | -95.9826
2.7109 | -95.8825
2.7146 | -95.7292
2.7183 | -95.5769
2.722 | -95.4682
2.7257 | -95.4137
2.7294 | -95.4076
2.7331 | -95.4306
2.7367 | -95.4618
2.7404 | -95.4892
2.7441 | -95.5261
2.7478 | -95.5909
2.7515 | -95.6794
2.7552 | -95.7762
2.7589 | -95.8581
2.7626 | -95.912
2.7663 | -95.9366
2.7699 | -95.9258
2.7736 | -95.8825
2.7773 | -95.8119
2.781 | -95.7462
2.7847 | -95.6927
2.7884 | -95.6659
2.7921 | -95.6535
2.7958 | -95.6513
2.7994 | -95.6887
2.8031 | -95.7634
2.8068 | -95.8518
2.8105 | -95.9223
2.8142 | -95.9419
2.8179 | -95.9317
2.8216 | -95.9291
2.8253 | -95.9703
2.829 | -96.0578
2.8326 | -96.1742
2.8363 | -96.2596
2.84 | -96.2866
2.8437 | -96.2574
2.8474 | -96.1798
2.8511 | -96.1042
2.8548 | -96.0475
2.8585 | -96.0157
2.8621 | -96.0063
2.8658 | -96.007
2.8695 | -96.021
2.8732 | -96.065
2.8769 | -96.1363
2.8806 | -96.2202
2.8843 | -96.2885
2.888 | -96.2963
2.8917 | -96.2344
2.8953 | -96.1354
2.899 | -96.0409
2.9027 | -95.9985
2.9064 | -96.0252
2.9101 | -96.1257
2.9138 | -96.2658
2.9175 | -96.4017
2.9212 | -96.4817
2.9249 | -96.507
2.9285 | -96.4938
2.9322 | -96.4693
2.9359 | -96.4682
2.9396 | -96.493
2.9433 | -96.5375
2.947 | -96.5675
2.9507 | -96.5479
2.9544 | -96.466
2.958 | -96.3168
2.9617 | -96.112
2.9654 | -95.8828
2.9691 | -95.6913
2.9728 | -95.5799
2.9765 | -95.5664
2.9802 | -95.64
2.9839 | -95.7698
2.9876 | -95.9263
2.9912 | -96.0949
2.9949 | -96.2747
2.9986 | -96.4525
3.0023 | -96.601
3.006 | -96.7201
3.0097 | -96.8157
3.0134 | -96.8996
3.0171 | -96.976
3.0207 | -97.0331
3.0244 | -97.0886
3.0281 | -97.1453
3.0318 | -97.1808
3.0355 | -97.1798
3.0392 | -97.1373
3.0429 | -97.0507
3.0466 | -96.9599
3.0503 | -96.8805
3.0539 | -96.8158
3.0576 | -96.7966
3.0613 | -96.8211
3.065 | -96.8852
3.0687 | -96.9969
3.0724 | -97.0996
3.0761 | -97.1496
3.0798 | -97.1282
3.0834 | -97.0351
3.0871 | -96.9006
3.0908 | -96.7869
3.0945 | -96.715
3.0982 | -96.7117
3.1019 | -96.7765
3.1056 | -96.867
3.1093 | -96.9713
3.113 | -97.0407
3.1166 | -97.033
3.1203 | -96.9611
3.124 | -96.857
3.1277 | -96.7775
3.1314 | -96.7522
3.1351 | -96.7878
3.1388 | -96.8747
3.1425 | -96.9988
3.1462 | -97.1387
3.1498 | -97.2691
3.1535 | -97.3402
3.1572 | -97.2975
3.1609 | -97.1771
3.1646 | -97.0145
3.1683 | -96.863
3.172 | -96.7637
3.1757 | -96.7364
3.1793 | -96.7977
3.183 | -96.9371
3.1867 | -97.102
3.1904 | -97.1428
3.1941 | -97.1826
3.1978 | -97.1671
3.2015 | -97.1088
3.2052 | -97.0493
3.2089 | -97.0016
3.2125 | -96.962
3.2162 | -96.9388
3.2199 | -96.9232
3.2236 | -96.9095
3.2273 | -96.8953
3.231 | -96.8882
3.2347 | -96.8898
3.2384 | -96.9038
3.242 | -96.9326
3.2457 | -96.9839
3.2494 | -97.0522
3.2531 | -97.1373
3.2568 | -97.228
3.2605 | -97.3068
3.2642 | -97.3672
3.2679 | -97.4005
3.2716 | -97.4147
3.2752 | -97.4117
3.2789 | -97.4013
3.2826 | -97.3821
3.2863 | -97.3438
3.29 | -97.2864
3.2937 | -97.2252
3.2974 | -97.1666
3.3011 | -97.1233
3.3047 | -97.1103
3.3084 | -97.1306
3.3121 | -97.1681
3.3158 | -97.2117
3.3195 | -97.2455
3.3232 | -97.2641
3.3269 | -97.267
3.3306 | -97.2584
3.3343 | -97.217
3.3379 | -97.1564
3.3416 | -97.0818
3.3453 | -96.991
3.349 | -96.9083
3.3527 | -96.8505
3.3564 | -96.8187
3.3601 | -96.8013
3.3638 | -96.8093
3.3674 | -96.8428
3.3711 | -96.8893
3.3748 | -96.9412
3.3785 | -96.9858
3.3822 | -97.0066
3.3859 | -97.0007
3.3896 | -96.9757
3.3933 | -96.9469
3.397 | -96.9357
3.4006 | -96.9677
3.4043 | -97.0139
3.408 | -97.0684
3.4117 | -97.1235
3.4154 | -97.1713
3.4191 | -97.1981
3.4228 | -97.2174
3.4265 | -97.2333
3.4302 | -97.2587
3.4338 | -97.3121
3.4375 | -97.4045
3.4412 | -97.531
3.4449 | -97.6783
3.4486 | -97.8147
3.4523 | -97.9357
3.456 | -98.0334
3.4597 | -98.107
3.4633 | -98.1465
3.467 | -98.1536
3.4707 | -98.1401
3.4744 | -98.1243
3.4781 | -98.1227
3.4818 | -98.1537
3.4855 | -98.1958
3.4892 | -98.2253
3.4929 | -98.2151
3.4965 | -98.1617
3.5002 | -98.0712
3.5039 | -97.9334
3.5076 | -97.772
3.5113 | -97.6224
3.515 | -97.518
3.5187 | -97.4704
3.5224 | -97.4993
3.526 | -97.5841
3.5297 | -97.6963
3.5334 | -97.7982
3.5371 | -97.8739
3.5408 | -97.9257
3.5445 | -97.9698
3.5482 | -98.0212
3.5519 | -98.0761
3.5556 | -98.1245
3.5592 | -98.1642
3.5629 | -98.1875
3.5666 | -98.1669
3.5703 | -98.0881
3.574 | -97.9682
3.5777 | -97.8166
3.5814 | -97.682
3.5851 | -97.5942
3.5888 | -97.5678
3.5924 | -97.5865
3.5961 | -97.6324
3.5998 | -97.676
3.6035 | -97.7091
3.6072 | -97.7243
3.6109 | -97.7291
3.6146 | -97.742
3.6183 | -97.7594
3.6219 | -97.8124
3.6256 | -97.8905
3.6293 | -97.9665
3.633 | -98.033
3.6367 | -98.064
3.6404 | -98.0499
3.6441 | -98.0051
3.6478 | -97.948
3.6515 | -97.9024
3.6551 | -97.8789
3.6588 | -97.8992
3.6625 | -97.9747
3.6662 | -98.072
3.6699 | -98.1616
3.6736 | -98.221
3.6773 | -98.2327
3.681 | -98.2034
3.6846 | -98.1566
3.6883 | -98.1134
3.692 | -98.0937
3.6957 | -98.1014
3.6994 | -98.1379
3.7031 | -98.2028
3.7068 | -98.2842
3.7105 | -98.3774
3.7142 | -98.4424
3.7178 | -98.4531
3.7215 | -98.3942
3.7252 | -98.3695
3.7289 | -98.3416
3.7326 | -98.3296
3.7363 | -98.3383
3.74 | -98.3751
3.7437 | -98.4218
3.7473 | -98.4556
3.751 | -98.4691
3.7547 | -98.4578
3.7584 | -98.4171
3.7621 | -98.3614
3.7658 | -98.3175
3.7695 | -98.284
3.7732 | -98.2634
3.7769 | -98.2648
3.7805 | -98.263
3.7842 | -98.2577
3.7879 | -98.2447
3.7916 | -98.2123
3.7953 | -98.1707
3.799 | -98.1298
3.8027 | -98.097
3.8064 | -98.0841
3.8101 | -98.103
3.8137 | -98.156
3.8174 | -98.2295
3.8211 | -98.3302
3.8248 | -98.4462
3.8285 | -98.5497
3.8322 | -98.6429
3.8359 | -98.7229
3.8396 | -98.7737
3.8432 | -98.8055
3.8469 | -98.8189
3.8506 | -98.8186
3.8543 | -98.8202
3.858 | -98.8176
3.8617 | -98.8134
3.8654 | -98.7979
3.8691 | -98.7682
3.8728 | -98.7261
3.8764 | -98.6611
3.8801 | -98.596
3.8838 | -98.5248
3.8875 | -98.463
3.8912 | -98.424
3.8949 | -98.3983
3.8986 | -98.3882
3.9023 | -98.3903
3.9059 | -98.3925
3.9096 | -98.3925
3.9133 | -98.3841
3.917 | -98.3763
3.9207 | -98.3744
3.9244 | -98.3798
3.9281 | -98.3991
3.9318 | -98.429
3.9355 | -98.4737
3.9391 | -98.5306
3.9428 | -98.5911
3.9465 | -98.657
3.9502 | -98.7165
3.9539 | -98.7675
3.9576 | -98.8143
3.9613 | -98.8509
3.965 | -98.8805
3.9686 | -98.9028
3.9723 | -98.9086
3.976 | -98.8849
3.9797 | -98.8164
3.9834 | -98.7178
3.9871 | -98.6148
3.9908 | -98.5152
3.9945 | -98.4406
3.9982 | -98.3923
4.0018 | -98.3775
4.0055 | -98.393
4.0092 | -98.4345
4.0129 | -98.507
4.0166 | -98.609
4.0203 | -98.7276
4.024 | -98.8477
4.0277 | -98.9577
4.0314 | -99.0429
4.035 | -99.0937
4.0387 | -99.1243
4.0424 | -99.1441
4.0461 | -99.1568
4.0498 | -99.1639
4.0535 | -99.1775
4.0572 | -99.1887
4.0609 | -99.1952
4.0645 | -99.1975
4.0682 | -99.1796
4.0719 | -99.1462
4.0756 | -99.101
4.0793 | -99.0444
4.083 | -98.9939
4.0867 | -98.9609
4.0904 | -98.9504
4.0941 | -98.9652
4.0977 | -99.0022
4.1014 | -99.0448
4.1051 | -99.0826
4.1088 | -99.1094
4.1125 | -99.1209
4.1162 | -99.1199
4.1199 | -99.116
4.1236 | -99.1264
4.1272 | -99.1448
4.1309 | -99.1718
4.1346 | -99.2014
4.1383 | -99.2281
4.142 | -99.2481
4.1457 | -99.2665
4.1494 | -99.2939
4.1531 | -99.3271
4.1568 | -99.3662
4.1604 | -99.4109
4.1641 | -99.4481
4.1678 | -99.4662
4.1715 | -99.4635
4.1752 | -99.4443
4.1789 | -99.4051
4.1826 | -99.3549
4.1863 | -99.306
4.1899 | -99.2692
4.1936 | -99.2476
4.1973 | -99.2234
4.201 | -99.1961
4.2047 | -99.1765
4.2084 | -99.1629
4.2121 | -99.1578
4.2158 | -99.178
4.2195 | -99.2261
4.2231 | -99.2956
4.2268 | -99.3822
4.2305 | -99.4694
4.2342 | -99.5447
4.2379 | -99.5997
4.2416 | -99.6361
4.2453 | -99.6485
4.249 | -99.6361
4.2527 | -99.6235
4.2563 | -99.6593
4.26 | -99.6834
4.2637 | -99.72
4.2674 | -99.7569
4.2711 | -99.777
4.2748 | -99.7806
4.2785 | -99.7595
4.2822 | -99.7018
4.2858 | -99.6101
4.2895 | -99.4964
4.2932 | -99.3749
4.2969 | -99.259
4.3006 | -99.1608
4.3043 | -99.0877
4.308 | -99.0412
4.3117 | -99.0223
4.3154 | -99.038
4.319 | -99.0823
4.3227 | -99.152
4.3264 | -99.2427
4.3301 | -99.3507
4.3338 | -99.4657
4.3375 | -99.5662
4.3412 | -99.651
4.3449 | -99.7039
4.3485 | -99.7256
4.3522 | -99.7155
4.3559 | -99.68
4.3596 | -99.6261
4.3633 | -99.5685
4.367 | -99.5222
4.3707 | -99.4995
4.3744 | -99.5034
4.3781 | -99.5287
4.3817 | -99.5711
4.3854 | -99.6157
4.3891 | -99.6597
4.3928 | -99.6982
4.3965 | -99.7306
4.4002 | -99.7534
4.4039 | -99.7638
4.4076 | -99.7677
4.4112 | -99.7541
4.4149 | -99.7216
4.4186 | -99.678
4.4223 | -99.6289
4.426 | -99.5672
4.4297 | -99.4939
4.4334 | -99.4161
4.4371 | -99.3348
4.4408 | -99.2507
4.4444 | -99.1685
4.4481 | -99.0938
4.4518 | -99.0335
4.4555 | -99.0021
4.4592 | -99.0026
4.4629 | -99.0367
4.4666 | -99.1065
4.4703 | -99.2095
4.474 | -99.3289
4.4776 | -99.4416
4.4813 | -99.5359
4.485 | -99.6089
4.4887 | -99.6507
4.4924 | -99.6778
4.4961 | -99.6934
4.4998 | -99.716
4.5035 | -99.7507
4.5071 | -99.8005
4.5108 | -99.8672
4.5145 | -99.9413
4.5182 | -100.0138
4.5219 | -100.082
4.5256 | -100.1258
4.5293 | -100.1611
4.533 | -100.1906
4.5367 | -100.2023
4.5403 | -100.1902
4.544 | -100.1702
4.5477 | -100.1416
4.5514 | -100.0972
4.5551 | -100.0387
4.5588 | -99.9705
4.5625 | -99.9048
4.5662 | -99.8579
4.5698 | -99.8216
4.5735 | -99.8007
4.5772 | -99.7989
4.5809 | -99.8188
4.5846 | -99.8635
4.5883 | -99.9243
4.592 | -100.0007
4.5957 | -100.0987
4.5994 | -100.1918
4.603 | -100.2726
4.6067 | -100.3424
4.6104 | -100.3856
4.6141 | -100.3908
4.6178 | -100.3751
4.6215 | -100.3485
4.6252 | -100.3226
4.6289 | -100.3053
4.6326 | -100.2829
4.6362 | -100.2569
4.6399 | -100.227
4.6436 | -100.2001
4.6473 | -100.1781
4.651 | -100.1509
4.6547 | -100.1272
4.6584 | -100.1116
4.6621 | -100.099
4.6657 | -100.0962
4.6694 | -100.1071
4.6731 | -100.1375
4.6768 | -100.1738
4.6805 | -100.2231
4.6842 | -100.2835
4.6879 | -100.3272
4.6916 | -100.3539
4.6953 | -100.3619
4.6989 | -100.3571
4.7026 | -100.3494
4.7063 | -100.3405
4.71 | -100.3399
4.7137 | -100.3297
4.7174 | -100.3257
4.7211 | -100.3207
4.7248 | -100.2949
4.7284 | -100.2472
4.7321 | -100.1707
4.7358 | -100.0775
4.7395 | -99.974
4.7432 | -99.8565
4.7469 | -99.7519
4.7506 | -99.6673
4.7543 | -99.6018
4.758 | -99.5802
4.7616 | -99.5993
4.7653 | -99.6561
4.769 | -99.7468
4.7727 | -99.8571
4.7764 | -99.9859
4.7801 | -100.1195
4.7838 | -100.2487
4.7875 | -100.3275
4.7911 | -100.4116
4.7948 | -100.4798
4.7985 | -100.54
4.8022 | -100.5826
4.8059 | -100.614
4.8096 | -100.627
4.8133 | -100.626
4.817 | -100.6173
4.8207 | -100.593
4.8243 | -100.5634
4.828 | -100.536
4.8317 | -100.504
4.8354 | -100.4715
4.8391 | -100.4423
4.8428 | -100.4199
4.8465 | -100.4048
4.8502 | -100.3921
4.8539 | -100.3814
4.8575 | -100.3803
4.8612 | -100.3851
4.8649 | -100.396
4.8686 | -100.4148
4.8723 | -100.4433
4.876 | -100.4747
4.8797 | -100.5225
4.8834 | -100.5845
4.887 | -100.6696
4.8907 | -100.767
4.8944 | -100.8682
4.8981 | -100.9597
4.9018 | -101.0246
4.9055 | -101.058
4.9092 | -101.0557
4.9129 | -101.0094
4.9166 | -100.9234
4.9202 | -100.8119
4.9239 | -100.6959
4.9276 | -100.5831
4.9313 | -100.4908
4.935 | -100.4306
4.9387 | -100.4053
4.9424 | -100.4133
4.9461 | -100.4602
4.9497 | -100.5377
4.9534 | -100.6341
4.9571 | -100.7319
4.9608 | -100.8252
4.9645 | -100.8942
4.9682 | -100.9422
4.9719 | -100.9579
4.9756 | -100.9417
4.9793 | -100.8922
4.9829 | -100.8192
4.9866 | -100.7361
4.9903 | -100.6598
4.994 | -100.6019
4.9977 | -100.5637
5.0014 | -100.5519
5.0051 | -100.567
5.0088 | -100.6141
5.0124 | -100.6794
5.0161 | -100.7602
5.0198 | -100.8493
5.0235 | -100.9326
5.0272 | -101.0037
5.0309 | -101.0666
5.0346 | -101.1243
5.0383 | -101.1646
5.042 | -101.1863
5.0456 | -101.2006
5.0493 | -101.2096
5.053 | -101.2082
5.0567 | -101.1927
5.0604 | -101.1513
5.0641 | -101.0895
5.0678 | -101.0134
5.0715 | -100.9218
5.0751 | -100.8139
5.0788 | -100.6963
5.0825 | -100.5893
5.0862 | -100.493
5.0899 | -100.4187
5.0936 | -100.3737
5.0973 | -100.3574
5.101 | -100.3703
5.1047 | -100.403
5.1083 | -100.4529
5.112 | -100.5103
5.1157 | -100.5729
5.1194 | -100.6305
5.1231 | -100.6708
5.1268 | -100.699
5.1305 | -100.7165
5.1342 | -100.7287
5.1379 | -100.7388
5.1415 | -100.745
5.1452 | -100.7531
5.1489 | -100.7608
5.1526 | -100.7747
5.1563 | -100.7825
5.16 | -100.7783
5.1637 | -100.7671
5.1674 | -100.7523
5.171 | -100.7344
5.1747 | -100.7131
5.1784 | -100.7005
5.1821 | -100.698
5.1858 | -100.7036
5.1895 | -100.7213
5.1932 | -100.7568
5.1969 | -100.7988
5.2006 | -100.8555
5.2042 | -100.9189
5.2079 | -100.9833
5.2116 | -101.0427
5.2153 | -101.0882
5.219 | -101.1182
5.2227 | -101.1358
5.2264 | -101.1497
5.2301 | -101.1571
5.2337 | -101.1721
5.2374 | -101.1857
5.2411 | -101.2057
5.2448 | -101.2257
5.2485 | -101.2391
5.2522 | -101.2576
5.2559 | -101.2671
5.2596 | -101.2672
5.2633 | -101.253
5.2669 | -101.2258
5.2706 | -101.1875
5.2743 | -101.147
5.278 | -101.1002
5.2817 | -101.0589
5.2854 | -101.0279
5.2891 | -101.011
5.2928 | -101.0096
5.2965 | -101.0181
5.3001 | -101.0239
5.3038 | -101.0292
5.3075 | -101.0387
5.3112 | -101.0427
5.3149 | -101.0528
5.3186 | -101.0738
5.3223 | -101.0861
5.326 | -101.1026
5.3296 | -101.1201
5.3333 | -101.1353
5.337 | -101.1471
5.3407 | -101.1519
5.3444 | -101.1527
5.3481 | -101.1504
5.3518 | -101.1493
5.3555 | -101.1512
5.3592 | -101.1611
5.3628 | -101.1869
5.3665 | -101.2296
5.3702 | -101.2906
5.3739 | -101.369
5.3776 | -101.4599
5.3813 | -101.56
5.385 | -101.6471
5.3887 | -101.7248
5.3923 | -101.7898
5.396 | -101.8287
5.3997 | -101.8435
5.4034 | -101.8404
5.4071 | -101.8121
5.4108 | -101.7562
5.4145 | -101.6783
5.4182 | -101.5904
5.4219 | -101.5008
5.4255 | -101.4205
5.4292 | -101.3481
5.4329 | -101.2918
5.4366 | -101.2487
5.4403 | -101.2334
5.444 | -101.2297
5.4477 | -101.231
5.4514 | -101.2444
5.455 | -101.2669
5.4587 | -101.2849
5.4624 | -101.3007
5.4661 | -101.3119
5.4698 | -101.3193
5.4735 | -101.3053
5.4772 | -101.2716
5.4809 | -101.227
5.4846 | -101.1693
5.4882 | -101.1028
5.4919 | -101.0373
5.4956 | -100.9857
5.4993 | -100.9481
5.503 | -100.9238
5.5067 | -100.9175
5.5104 | -100.9244
5.5141 | -100.9451
5.5178 | -100.9718
5.5214 | -100.9998
5.5251 | -101.0169
5.5288 | -101.0277
5.5325 | -101.0253
5.5362 | -101.0078
5.5399 | -100.9824
5.5436 | -100.9575
5.5473 | -100.9295
5.5509 | -100.9031
5.5546 | -100.8818
5.5583 | -100.88
5.562 | -100.9009
5.5657 | -100.9447
5.5694 | -101.0116
5.5731 | -101.0958
5.5768 | -101.1847
5.5805 | -101.2709
5.5841 | -101.3596
5.5878 | -101.4435
5.5915 | -101.5164
5.5952 | -101.5764
5.5989 | -101.6143
5.6026 | -101.6292
5.6063 | -101.6256
5.61 | -101.604
5.6136 | -101.5664
5.6173 | -101.5217
5.621 | -101.4641
5.6247 | -101.4005
5.6284 | -101.344
5.6321 | -101.3058
5.6358 | -101.2841
5.6395 | -101.2784
5.6432 | -101.284
5.6468 | -101.3013
5.6505 | -101.3198
5.6542 | -101.3443
5.6579 | -101.3768
5.6616 | -101.4164
5.6653 | -101.4528
5.669 | -101.4756
5.6727 | -101.4912
5.6763 | -101.5029
5.68 | -101.5087
5.6837 | -101.5108
5.6874 | -101.51
5.6911 | -101.4944
5.6948 | -101.4755
5.6985 | -101.4493
5.7022 | -101.4227
5.7059 | -101.3976
5.7095 | -101.3706
5.7132 | -101.3393
5.7169 | -101.3065
5.7206 | -101.273
5.7243 | -101.2441
5.728 | -101.221
5.7317 | -101.2015
5.7354 | -101.1941
5.739 | -101.1901
5.7427 | -101.1924
5.7464 | -101.2058
5.7501 | -101.2333
5.7538 | -101.2797
5.7575 | -101.339
5.7612 | -101.4061
5.7649 | -101.4814
5.7686 | -101.5637
5.7722 | -101.6501
5.7759 | -101.7342
5.7796 | -101.7981
5.7833 | -101.8535
5.787 | -101.8832
5.7907 | -101.9017
5.7944 | -101.9031
5.7981 | -101.8882
5.8018 | -101.8537
5.8054 | -101.8146
5.8091 | -101.7715
5.8128 | -101.7321
5.8165 | -101.7062
5.8202 | -101.6956
5.8239 | -101.6981
5.8276 | -101.7113
5.8313 | -101.733
5.8349 | -101.7676
5.8386 | -101.8003
5.8423 | -101.826
5.846 | -101.8431
5.8497 | -101.8303
5.8534 | -101.8259
5.8571 | -101.8059
5.8608 | -101.7777
5.8645 | -101.7543
5.8681 | -101.7281
5.8718 | -101.6982
5.8755 | -101.6719
5.8792 | -101.6368
5.8829 | -101.6084
5.8866 | -101.5887
5.8903 | -101.5737
5.894 | -101.5638
5.8976 | -101.5569
5.9013 | -101.5541
5.905 | -101.5573
5.9087 | -101.5622
5.9124 | -101.5722
5.9161 | -101.5877
5.9198 | -101.6081
5.9235 | -101.6261
5.9272 | -101.6514
5.9308 | -101.6848
5.9345 | -101.7183
5.9382 | -101.7505
5.9419 | -101.7835
5.9456 | -101.818
5.9493 | -101.8553
5.953 | -101.8952
5.9567 | -101.9259
5.9604 | -101.9619
5.964 | -101.9928
5.9677 | -102.0256
5.9714 | -102.0629
5.9751 | -102.0837
5.9788 | -102.0937
5.9825 | -102.0985
5.9862 | -102.0981
5.9899 | -102.0945
5.9935 | -102.0855
5.9972 | -102.0714
6.0009 | -102.0594
6.0046 | -102.0474
6.0083 | -102.0322
6.012 | -102.013
6.0157 | -101.9923
6.0194 | -101.972
6.0231 | -101.9473
6.0267 | -101.9241
6.0304 | -101.8992
6.0341 | -101.883
6.0378 | -101.877
6.0415 | -101.8775
6.0452 | -101.8855
6.0489 | -101.9084
6.0526 | -101.935
6.0562 | -101.9709
6.0599 | -102.0121
6.0636 | -102.0479
6.0673 | -102.0939
6.071 | -102.1356
6.0747 | -102.1631
6.0784 | -102.1807
6.0821 | -102.1871
6.0858 | -102.1857
6.0894 | -102.1741
6.0931 | -102.1406
6.0968 | -102.0932
6.1005 | -102.0325
6.1042 | -101.9663
6.1079 | -101.8972
6.1116 | -101.8269
6.1153 | -101.7619
6.1189 | -101.7068
6.1226 | -101.663
6.1263 | -101.6289
6.13 | -101.6062
6.1337 | -101.5999
6.1374 | -101.6069
6.1411 | -101.6224
6.1448 | -101.65
6.1485 | -101.6867
6.1521 | -101.7272
6.1558 | -101.7762
6.1595 | -101.8235
6.1632 | -101.8668
6.1669 | -101.9088
6.1706 | -101.9383
6.1743 | -101.9505
6.178 | -101.9445
6.1817 | -101.9338
6.1853 | -101.9134
6.189 | -101.8855
6.1927 | -101.8548
6.1964 | -101.8219
6.2001 | -101.7918
6.2038 | -101.7649
6.2075 | -101.7499
6.2112 | -101.744
6.2148 | -101.7503
6.2185 | -101.7695
6.2222 | -101.7978
6.2259 | -101.8326
6.2296 | -101.8726
6.2333 | -101.9195
6.237 | -101.9665
6.2407 | -102.004
6.2444 | -102.0337
6.248 | -102.0574
6.2517 | -102.0784
6.2554 | -102.095
6.2591 | -102.1061
6.2628 | -102.1154
6.2665 | -102.129
6.2702 | -102.1414
6.2739 | -102.1449
6.2775 | -102.1529
6.2812 | -102.1596
6.2849 | -102.1686
6.2886 | -102.175
6.2923 | -102.1857
6.296 | -102.2068
6.2997 | -102.2298
6.3034 | -102.2511
6.3071 | -102.2779
6.3107 | -102.3121
6.3144 | -102.3483
6.3181 | -102.3894
6.3218 | -102.4233
6.3255 | -102.4498
6.3292 | -102.4688
6.3329 | -102.4871
6.3366 | -102.499
6.3402 | -102.502
6.3439 | -102.5008
6.3476 | -102.5015
6.3513 | -102.5
6.355 | -102.4965
6.3587 | -102.4941
6.3624 | -102.5001
6.3661 | -102.5133
6.3698 | -102.5263
6.3734 | -102.5372
6.3771 | -102.5491
6.3808 | -102.5661
6.3845 | -102.5747
6.3882 | -102.5849
6.3919 | -102.5912
6.3956 | -102.6012
6.3993 | -102.617
6.403 | -102.6363
6.4066 | -102.6589
6.4103 | -102.6854
6.414 | -102.7138
6.4177 | -102.7502
6.4214 | -102.7859
6.4251 | -102.8185
6.4288 | -102.8462
6.4325 | -102.8697
6.4361 | -102.8881
6.4398 | -102.9004
6.4435 | -102.9122
6.4472 | -102.9171
6.4509 | -102.9053
6.4546 | -102.8776
6.4583 | -102.8355
6.462 | -102.7803
6.4657 | -102.7192
6.4693 | -102.6549
6.473 | -102.5801
6.4767 | -102.4963
6.4804 | -102.4175
6.4841 | -102.3422
6.4878 | -102.2734
6.4915 | -102.2116
6.4952 | -102.1581
6.4988 | -102.1222
6.5025 | -102.1021
6.5062 | -102.0982
6.5099 | -102.1181
6.5136 | -102.1545
6.5173 | -102.2034
6.521 | -102.2647
6.5247 | -102.3307
6.5284 | -102.397
6.532 | -102.4526
6.5357 | -102.4944
6.5394 | -102.5134
6.5431 | -102.5158
6.5468 | -102.5085
6.5505 | -102.4939
6.5542 | -102.4725
6.5579 | -102.4454
6.5615 | -102.4192
6.5652 | -102.3971
6.5689 | -102.3799
6.5726 | -102.3757
6.5763 | -102.3836
6.58 | -102.3989
6.5837 | -102.426
6.5874 | -102.4602
6.5911 | -102.4951
6.5947 | -102.5236
6.5984 | -102.5417
6.6021 | -102.5501
6.6058 | -102.5503
6.6095 | -102.5365
6.6132 | -102.5126
6.6169 | -102.4822
6.6206 | -102.4395
6.6243 | -102.3918
6.6279 | -102.3374
6.6316 | -102.2913
6.6353 | -102.2504
6.639 | -102.2211
6.6427 | -102.1901
6.6464 | -102.1558
6.6501 | -102.1296
6.6538 | -102.1062
6.6574 | -102.0819
6.6611 | -102.0536
6.6648 | -102.0227
6.6685 | -101.9988
6.6722 | -101.9782
6.6759 | -101.9628
6.6796 | -101.9515
6.6833 | -101.9516
6.687 | -101.9668
6.6906 | -101.9991
6.6943 | -102.0437
6.698 | -102.0948
6.7017 | -102.1467
6.7054 | -102.2012
6.7091 | -102.255
6.7128 | -102.3042
6.7165 | -102.3532
6.7201 | -102.4046
6.7238 | -102.4543
6.7275 | -102.5061
6.7312 | -102.5493
6.7349 | -102.589
6.7386 | -102.632
6.7423 | -102.6791
6.746 | -102.7363
6.7497 | -102.7809
6.7533 | -102.8239
6.757 | -102.8677
6.7607 | -102.9097
6.7644 | -102.9391
6.7681 | -102.9556
6.7718 | -102.9609
6.7755 | -102.9542
6.7792 | -102.9371
6.7828 | -102.9064
6.7865 | -102.8598
6.7902 | -102.8066
6.7939 | -102.753
6.7976 | -102.6936
6.8013 | -102.6417
6.805 | -102.5899
6.8087 | -102.5404
6.8124 | -102.5004
6.816 | -102.4642
6.8197 | -102.4297
6.8234 | -102.4002
6.8271 | -102.3752
6.8308 | -102.3552
6.8345 | -102.3411
6.8382 | -102.327
6.8419 | -102.3144
6.8456 | -102.3036
6.8492 | -102.3028
6.8529 | -102.3012
6.8566 | -102.3034
6.8603 | -102.3123
6.864 | -102.3258
6.8677 | -102.3345
6.8714 | -102.3469
6.8751 | -102.3674
6.8787 | -102.3904
6.8824 | -102.4184
6.8861 | -102.4549
6.8898 | -102.4951
6.8935 | -102.5359
6.8972 | -102.5936
6.9009 | -102.6539
6.9046 | -102.7212
6.9083 | -102.79
6.9119 | -102.8588
6.9156 | -102.9189
6.9193 | -102.9813
6.923 | -103.0333
6.9267 | -103.0734
6.9304 | -103.0986
6.9341 | -103.1067
6.9378 | -103.0998
6.9414 | -103.08
6.9451 | -103.0509
6.9488 | -102.9995
6.9525 | -102.9398
6.9562 | -102.8799
6.9599 | -102.8173
6.9636 | -102.7513
6.9673 | -102.6916
6.971 | -102.6434
6.9746 | -102.6066
6.9783 | -102.5758
6.982 | -102.5588
6.9857 | -102.5544
6.9894 | -102.5642
6.9931 | -102.5864
6.9968 | -102.6245
7.0005 | -102.6604
7.0041 | -102.707
7.0078 | -102.7577
7.0115 | -102.8135
7.0152 | -102.8664
7.0189 | -102.9162
7.0226 | -102.9632
7.0263 | -103.0068
7.03 | -103.0598
7.0337 | -103.1132
7.0373 | -103.1683
7.041 | -103.2235
7.0447 | -103.2689
7.0484 | -103.3069
7.0521 | -103.3476
7.0558 | -103.3865
7.0595 | -103.4142
7.0632 | -103.4332
7.0669 | -103.4431
7.0705 | -103.4453
7.0742 | -103.4363
7.0779 | -103.4172
7.0816 | -103.3928
7.0853 | -103.3657
7.089 | -103.334
7.0927 | -103.3051
7.0964 | -103.2774
7.1 | -103.2541
7.1037 | -103.2387
7.1074 | -103.2326
7.1111 | -103.2291
7.1148 | -103.2377
7.1185 | -103.2512
7.1222 | -103.2658
7.1259 | -103.286
7.1296 | -103.308
7.1332 | -103.3229
7.1369 | -103.3271
7.1406 | -103.3318
7.1443 | -103.3275
7.148 | -103.3187
7.1517 | -103.31
7.1554 | -103.2976
7.1591 | -103.277
7.1627 | -103.2583
7.1664 | -103.24
7.1701 | -103.2204
7.1738 | -103.208
7.1775 | -103.1909
7.1812 | -103.1736
7.1849 | -103.154
7.1886 | -103.1281
7.1923 | -103.0989
7.1959 | -103.0726
7.1996 | -103.0391
7.2033 | -103.0011
7.207 | -102.9636
7.2107 | -102.9246
7.2144 | -102.8877
7.2181 | -102.8562
7.2218 | -102.8312
7.2255 | -102.8136
7.2291 | -102.8034
7.2328 | -102.7996
7.2365 | -102.8059
7.2402 | -102.8228
7.2439 | -102.8464
7.2476 | -102.8701
7.2513 | -102.8997
7.255 | -102.9312
7.2586 | -102.9638
7.2623 | -102.9958
7.266 | -103.0251
7.2697 | -103.0469
7.2734 | -103.0662
7.2771 | -103.0825
7.2808 | -103.0909
7.2845 | -103.1001
7.2882 | -103.1025
7.2918 | -103.1025
7.2955 | -103.0965
7.2992 | -103.086
7.3029 | -103.07
7.3066 | -103.0568
7.3103 | -103.0479
7.314 | -103.0371
7.3177 | -103.0301
7.3213 | -103.0286
7.325 | -103.0231
7.3287 | -103.0168
7.3324 | -103.0184
7.3361 | -103.0209
7.3398 | -103.0258
7.3435 | -103.0293
7.3472 | -103.032
7.3509 | -103.0315
7.3545 | -103.0311
7.3582 | -103.0246
7.3619 | -103.0154
7.3656 | -103.0028
7.3693 | -102.9827
7.373 | -102.9591
7.3767 | -102.9362
7.3804 | -102.9123
7.384 | -102.8881
7.3877 | -102.8653
7.3914 | -102.8486
7.3951 | -102.8386
7.3988 | -102.8341
7.4025 | -102.8391
7.4062 | -102.857
7.4099 | -102.8806
7.4136 | -102.9123
7.4172 | -102.9509
7.4209 | -102.9932
7.4246 | -103.0395
7.4283 | -103.0961
7.432 | -103.1578
7.4357 | -103.2194
7.4394 | -103.2782
7.4431 | -103.3323
7.4467 | -103.3697
7.4504 | -103.4072
7.4541 | -103.4446
7.4578 | -103.4779
7.4615 | -103.502
7.4652 | -103.5215
7.4689 | -103.5388
7.4726 | -103.547
7.4763 | -103.5517
7.4799 | -103.5496
7.4836 | -103.5378
7.4873 | -103.5239
7.491 | -103.5063
7.4947 | -103.4803
7.4984 | -103.4557
7.5021 | -103.4269
7.5058 | -103.3939
7.5095 | -103.3646
7.5131 | -103.3413
7.5168 | -103.3171
7.5205 | -103.3013
7.5242 | -103.2896
7.5279 | -103.2775
7.5316 | -103.271
7.5353 | -103.2652
7.539 | -103.2578
7.5426 | -103.2567
7.5463 | -103.2598
7.55 | -103.2617
7.5537 | -103.2663
7.5574 | -103.2646
7.5611 | -103.2636
7.5648 | -103.2632
7.5685 | -103.2594
7.5722 | -103.2548
7.5758 | -103.2515
7.5795 | -103.2512
7.5832 | -103.2583
7.5869 | -103.2727
7.5906 | -103.2963
7.5943 | -103.3344
7.598 | -103.382
7.6017 | -103.431
7.6053 | -103.4844
7.609 | -103.5516
7.6127 | -103.6174
7.6164 | -103.6815
7.6201 | -103.7443
7.6238 | -103.8003
7.6275 | -103.8468
7.6312 | -103.8852
7.6349 | -103.9192
7.6385 | -103.9463
7.6422 | -103.9665
7.6459 | -103.9776
7.6496 | -103.9713
7.6533 | -103.9516
7.657 | -103.9199
7.6607 | -103.8793
7.6644 | -103.8356
7.6681 | -103.7864
7.6717 | -103.7327
7.6754 | -103.6734
7.6791 | -103.6179
7.6828 | -103.562
7.6865 | -103.5153
7.6902 | -103.4744
7.6939 | -103.4311
7.6976 | -103.3947
7.7012 | -103.3697
7.7049 | -103.3485
7.7086 | -103.3335
7.7123 | -103.3245
7.716 | -103.3153
7.7197 | -103.3091
7.7234 | -103.3027
7.7271 | -103.3028
7.7308 | -103.3088
7.7344 | -103.3173
7.7381 | -103.3281
7.7418 | -103.3463
7.7455 | -103.3665
7.7492 | -103.3752
7.7529 | -103.3925
7.7566 | -103.4173
7.7603 | -103.4391
7.7639 | -103.4572
7.7676 | -103.4721
7.7713 | -103.4856
7.775 | -103.4925
7.7787 | -103.4921
7.7824 | -103.4882
7.7861 | -103.48
7.7898 | -103.4657
7.7935 | -103.4479
7.7971 | -103.4316
7.8008 | -103.4109
7.8045 | -103.3909
7.8082 | -103.3783
7.8119 | -103.3689
7.8156 | -103.3599
7.8193 | -103.3548
7.823 | -103.3507
7.8266 | -103.3464
7.8303 | -103.3466
7.834 | -103.346
7.8377 | -103.3453
7.8414 | -103.3444
7.8451 | -103.3409
7.8488 | -103.339
7.8525 | -103.3366
7.8562 | -103.3341
7.8598 | -103.3376
7.8635 | -103.3464
7.8672 | -103.3449
7.8709 | -103.3378
7.8746 | -103.3343
7.8783 | -103.3322
7.882 | -103.3234
7.8857 | -103.3106
7.8894 | -103.2906
7.893 | -103.2651
7.8967 | -103.2392
7.9004 | -103.2187
7.9041 | -103.2012
7.9078 | -103.182
7.9115 | -103.1633
7.9152 | -103.1479
7.9189 | -103.1363
7.9225 | -103.1223
7.9262 | -103.109
7.9299 | -103.1042
7.9336 | -103.1088
7.9373 | -103.1229
7.941 | -103.147
7.9447 | -103.1802
7.9484 | -103.2212
7.9521 | -103.2708
7.9557 | -103.3275
7.9594 | -103.3935
7.9631 | -103.4615
7.9668 | -103.5277
7.9705 | -103.5886
7.9742 | -103.6452
7.9779 | -103.6946
7.9816 | -103.7056
7.9852 | -103.7336
7.9889 | -103.7502
7.9926 | -103.758
7.9963 | -103.7614
8 | -103.7589
8.0037 | -103.7507
8.0074 | -103.744
8.0111 | -103.7432
8.0148 | -103.7423
8.0184 | -103.7386
8.0221 | -103.733
8.0258 | -103.7339
8.0295 | -103.7413
8.0332 | -103.7543
8.0369 | -103.7695
8.0406 | -103.789
8.0443 | -103.8093
8.0479 | -103.8222
8.0516 | -103.8357
8.0553 | -103.8551
8.059 | -103.8735
8.0627 | -103.8885
8.0664 | -103.9008
8.0701 | -103.9109
8.0738 | -103.9144
8.0775 | -103.9123
8.0811 | -103.9071
8.0848 | -103.8935
8.0885 | -103.8762
8.0922 | -103.8506
8.0959 | -103.8245
8.0996 | -103.7981
8.1033 | -103.7669
8.107 | -103.7357
8.1106 | -103.7076
8.1143 | -103.6788
8.118 | -103.652
8.1217 | -103.6272
8.1254 | -103.5988
8.1291 | -103.572
8.1328 | -103.5462
8.1365 | -103.5271
8.1402 | -103.5084
8.1438 | -103.4964
8.1475 | -103.4818
8.1512 | -103.4725
8.1549 | -103.4645
8.1586 | -103.4621
8.1623 | -103.4655
8.166 | -103.4726
8.1697 | -103.4857
8.1734 | -103.499
8.177 | -103.5122
8.1807 | -103.5305
8.1844 | -103.5551
8.1881 | -103.5848
8.1918 | -103.6123
8.1955 | -103.6418
8.1992 | -103.6782
8.2029 | -103.7183
8.2065 | -103.7623
8.2102 | -103.8006
8.2139 | -103.8358
8.2176 | -103.8688
8.2213 | -103.9004
8.225 | -103.9305
8.2287 | -103.9557
8.2324 | -103.9721
8.2361 | -103.9841
8.2397 | -103.99
8.2434 | -103.9927
8.2471 | -103.9924
8.2508 | -103.9829
8.2545 | -103.9673
8.2582 | -103.9466
8.2619 | -103.929
8.2656 | -103.9072
8.2692 | -103.8843
8.2729 | -103.8533
8.2766 | -103.8241
8.2803 | -103.7944
8.284 | -103.7624
8.2877 | -103.7308
8.2914 | -103.6999
8.2951 | -103.6692
8.2988 | -103.6405
8.3024 | -103.615
8.3061 | -103.5938
8.3098 | -103.5776
8.3135 | -103.567
8.3172 | -103.5592
8.3209 | -103.5514
8.3246 | -103.5495
8.3283 | -103.5567
8.332 | -103.5754
8.3356 | -103.5998
8.3393 | -103.6278
8.343 | -103.6636
8.3467 | -103.7016
8.3504 | -103.7418
8.3541 | -103.7864
8.3578 | -103.8308
8.3615 | -103.8754
8.3651 | -103.9218
8.3688 | -103.9664
8.3725 | -104.0056
8.3762 | -104.0347
8.3799 | -104.0537
8.3836 | -104.0728
8.3873 | -104.0897
8.391 | -104.0984
8.3947 | -104.102
8.3983 | -104.0962
8.402 | -104.0848
8.4057 | -104.0764
8.4094 | -104.0732
8.4131 | -104.0687
8.4168 | -104.0613
8.4205 | -104.0497
8.4242 | -104.0331
8.4278 | -104.0148
8.4315 | -103.9963
8.4352 | -103.9769
8.4389 | -103.9514
8.4426 | -103.9192
8.4463 | -103.8864
8.45 | -103.8538
8.4537 | -103.8206
8.4574 | -103.7919
8.461 | -103.7703
8.4647 | -103.7528
8.4684 | -103.7457
8.4721 | -103.7488
8.4758 | -103.7594
8.4795 | -103.774
8.4832 | -103.7941
8.4869 | -103.8137
8.4905 | -103.8399
8.4942 | -103.8757
8.4979 | -103.912
8.5016 | -103.9492
8.5053 | -103.9839
8.509 | -104.0138
8.5127 | -104.0353
8.5164 | -104.0599
8.5201 | -104.0822
8.5237 | -104.1047
8.5274 | -104.1197
8.5311 | -104.1294
8.5348 | -104.1304
8.5385 | -104.1314
8.5422 | -104.1299
8.5459 | -104.1282
8.5496 | -104.127
8.5533 | -104.1181
8.5569 | -104.1077
8.5606 | -104.1056
8.5643 | -104.1062
8.568 | -104.1113
8.5717 | -104.1114
8.5754 | -104.1114
8.5791 | -104.1131
8.5828 | -104.1145
8.5864 | -104.1141
8.5901 | -104.1111
8.5938 | -104.1032
8.5975 | -104.0928
8.6012 | -104.08
8.6049 | -104.0725
8.6086 | -104.0654
8.6123 | -104.0611
8.616 | -104.0489
8.6196 | -104.0319
8.6233 | -104.0239
8.627 | -104.0154
8.6307 | -104.0038
8.6344 | -103.9963
8.6381 | -103.986
8.6418 | -103.9717
8.6455 | -103.9614
8.6491 | -103.9466
8.6528 | -103.9304
8.6565 | -103.9113
8.6602 | -103.8899
8.6639 | -103.8718
8.6676 | -103.8523
8.6713 | -103.8339
8.675 | -103.8137
8.6787 | -103.7957
8.6823 | -103.7829
8.686 | -103.7694
8.6897 | -103.755
8.6934 | -103.7472
8.6971 | -103.7416
8.7008 | -103.743
8.7045 | -103.748
8.7082 | -103.7564
8.7118 | -103.7716
8.7155 | -103.7926
8.7192 | -103.8228
8.7229 | -103.861
8.7266 | -103.905
8.7303 | -103.9537
8.734 | -104.0052
8.7377 | -104.0616
8.7414 | -104.1223
8.745 | -104.1906
8.7487 | -104.2594
8.7524 | -104.3213
8.7561 | -104.3842
8.7598 | -104.4451
8.7635 | -104.4967
8.7672 | -104.5365
8.7709 | -104.5698
8.7745 | -104.5972
8.7782 | -104.614
8.7819 | -104.6223
8.7856 | -104.6261
8.7893 | -104.6159
8.793 | -104.6029
8.7967 | -104.5861
8.8004 | -104.5664
8.8041 | -104.546
8.8077 | -104.5207
8.8114 | -104.4943
8.8151 | -104.4685
8.8188 | -104.4416
8.8225 | -104.4191
8.8262 | -104.3996
8.8299 | -104.3813
8.8336 | -104.3626
8.8373 | -104.3457
8.8409 | -104.3318
8.8446 | -104.3197
8.8483 | -104.3095
8.852 | -104.2996
8.8557 | -104.2909
8.8594 | -104.285
8.8631 | -104.2767
8.8668 | -104.2682
8.8704 | -104.2632
8.8741 | -104.2575
8.8778 | -104.2498
8.8815 | -104.2439
8.8852 | -104.2383
8.8889 | -104.2352
8.8926 | -104.2427
8.8963 | -104.2569
8.9 | -104.2748
8.9036 | -104.2892
8.9073 | -104.3042
8.911 | -104.3227
8.9147 | -104.3449
8.9184 | -104.3705
8.9221 | -104.3884
8.9258 | -104.4081
8.9295 | -104.4221
8.9331 | -104.4323
8.9368 | -104.4413
8.9405 | -104.4474
8.9442 | -104.4498
8.9479 | -104.4512
8.9516 | -104.4457
8.9553 | -104.438
8.959 | -104.4298
8.9627 | -104.4236
8.9663 | -104.4176
8.97 | -104.4071
8.9737 | -104.3966
8.9774 | -104.3876
8.9811 | -104.3793
8.9848 | -104.3774
8.9885 | -104.3761
8.9922 | -104.3787
8.9959 | -104.3817
8.9995 | -104.3773
9.0032 | -104.3744
9.0069 | -104.3783
9.0106 | -104.3842
9.0143 | -104.3831
9.018 | -104.3734
9.0217 | -104.3606
9.0254 | -104.3502
9.029 | -104.3367
9.0327 | -104.3257
9.0364 | -104.3168
9.0401 | -104.302
9.0438 | -104.2824
9.0475 | -104.2674
9.0512 | -104.2519
9.0549 | -104.2397
9.0586 | -104.2254
9.0622 | -104.2118
9.0659 | -104.1955
9.0696 | -104.1828
9.0733 | -104.1751
9.077 | -104.1688
9.0807 | -104.1647
9.0844 | -104.1674
9.0881 | -104.1707
9.0917 | -104.1787
9.0954 | -104.1853
9.0991 | -104.1878
9.1028 | -104.1917
9.1065 | -104.1933
9.1102 | -104.1937
9.1139 | -104.1976
9.1176 | -104.1956
9.1213 | -104.1919
9.1249 | -104.1857
9.1286 | -104.1788
9.1323 | -104.1711
9.136 | -104.1582
9.1397 | -104.1399
9.1434 | -104.1199
9.1471 | -104.1019
9.1508 | -104.0859
9.1544 | -104.0704
9.1581 | -104.0516
9.1618 | -104.0366
9.1655 | -104.0239
9.1692 | -104.0126
9.1729 | -104.007
9.1766 | -104.0107
9.1803 | -104.0176
9.184 | -104.0239
9.1876 | -104.0397
9.1913 | -104.0643
9.195 | -104.0943
9.1987 | -104.1292
9.2024 | -104.1714
9.2061 | -104.2123
9.2098 | -104.2509
9.2135 | -104.295
9.2172 | -104.3383
9.2208 | -104.3792
9.2245 | -104.4237
9.2282 | -104.4635
9.2319 | -104.5011
9.2356 | -104.5398
9.2393 | -104.5748
9.243 | -104.6058
9.2467 | -104.6305
9.2503 | -104.6486
9.254 | -104.6647
9.2577 | -104.6783
9.2614 | -104.69
9.2651 | -104.701
9.2688 | -104.7105
9.2725 | -104.7208
9.2762 | -104.7288
9.2799 | -104.7357
9.2835 | -104.7402
9.2872 | -104.7463
9.2909 | -104.7501
9.2946 | -104.7538
9.2983 | -104.7618
9.302 | -104.7729
9.3057 | -104.7853
9.3094 | -104.7979
9.313 | -104.8062
9.3167 | -104.819
9.3204 | -104.8347
9.3241 | -104.842
9.3278 | -104.848
9.3315 | -104.8576
9.3352 | -104.8678
9.3389 | -104.8709
9.3426 | -104.8725
9.3462 | -104.8778
9.3499 | -104.8746
9.3536 | -104.8743
9.3573 | -104.8726
9.361 | -104.8636
9.3647 | -104.8541
9.3684 | -104.8425
9.3721 | -104.8303
9.3757 | -104.8152
9.3794 | -104.7974
9.3831 | -104.7844
9.3868 | -104.7721
9.3905 | -104.7603
9.3942 | -104.7515
9.3979 | -104.7422
9.4016 | -104.7344
9.4053 | -104.726
9.4089 | -104.7189
9.4126 | -104.7118
9.4163 | -104.7031
9.42 | -104.6985
9.4237 | -104.69
9.4274 | -104.6783
9.4311 | -104.6672
9.4348 | -104.651
9.4385 | -104.6305
9.4421 | -104.6119
9.4458 | -104.596
9.4495 | -104.5797
9.4532 | -104.5672
9.4569 | -104.5572
9.4606 | -104.5519
9.4643 | -104.5476
9.468 | -104.5452
9.4716 | -104.5501
9.4753 | -104.5597
9.479 | -104.5727
9.4827 | -104.5882
9.4864 | -104.611
9.4901 | -104.6314
9.4938 | -104.6523
9.4975 | -104.679
9.5012 | -104.7051
9.5048 | -104.7349
9.5085 | -104.76
9.5122 | -104.7823
9.5159 | -104.8033
9.5196 | -104.8151
9.5233 | -104.8253
9.527 | -104.8307
9.5307 | -104.8261
9.5343 | -104.8208
9.538 | -104.8082
9.5417 | -104.7977
9.5454 | -104.784
9.5491 | -104.7679
9.5528 | -104.7551
9.5565 | -104.7454
9.5602 | -104.7383
9.5639 | -104.7304
9.5675 | -104.7291
9.5712 | -104.7339
9.5749 | -104.7485
9.5786 | -104.7589
9.5823 | -104.7699
9.586 | -104.7804
9.5897 | -104.7938
9.5934 | -104.8031
9.597 | -104.8129
9.6007 | -104.8218
9.6044 | -104.8271
9.6081 | -104.8279
9.6118 | -104.8276
9.6155 | -104.8263
9.6192 | -104.82
9.6229 | -104.8089
9.6266 | -104.797
9.6302 | -104.7838
9.6339 | -104.7659
9.6376 | -104.7407
9.6413 | -104.7132
9.645 | -104.6836
9.6487 | -104.6528
9.6524 | -104.6214
9.6561 | -104.5916
9.6598 | -104.5636
9.6634 | -104.5351
9.6671 | -104.5076
9.6708 | -104.4848
9.6745 | -104.4598
9.6782 | -104.4374
9.6819 | -104.4156
9.6856 | -104.3965
9.6893 | -104.384
9.6929 | -104.3702
9.6966 | -104.3616
9.7003 | -104.3574
9.704 | -104.3592
9.7077 | -104.3602
9.7114 | -104.3602
9.7151 | -104.3644
9.7188 | -104.3721
9.7225 | -104.3833
9.7261 | -104.3931
9.7298 | -104.3991
9.7335 | -104.406
9.7372 | -104.4136
9.7409 | -104.4216
9.7446 | -104.4309
9.7483 | -104.4405
9.752 | -104.4523
9.7556 | -104.458
9.7593 | -104.462
9.763 | -104.4682
9.7667 | -104.4723
9.7704 | -104.4796
9.7741 | -104.4872
9.7778 | -104.4964
9.7815 | -104.5062
9.7852 | -104.518
9.7888 | -104.5289
9.7925 | -104.5419
9.7962 | -104.5547
9.7999 | -104.5664
9.8036 | -104.5779
9.8073 | -104.5875
9.811 | -104.598
9.8147 | -104.608
9.8183 | -104.6191
9.822 | -104.629
9.8257 | -104.6397
9.8294 | -104.6508
9.8331 | -104.6596
9.8368 | -104.6709
9.8405 | -104.6799
9.8442 | -104.6901
9.8479 | -104.7025
9.8515 | -104.7173
9.8552 | -104.7297
9.8589 | -104.7424
9.8626 | -104.7547
9.8663 | -104.7652
9.87 | -104.7775
9.8737 | -104.792
9.8774 | -104.8056
9.8811 | -104.818
9.8847 | -104.8277
9.8884 | -104.8372
9.8921 | -104.844
9.8958 | -104.8481
9.8995 | -104.8552
9.9032 | -104.8598
9.9069 | -104.8608
9.9106 | -104.8629
9.9142 | -104.8557
9.9179 | -104.8468
9.9216 | -104.8339
9.9253 | -104.8226
9.929 | -104.8085
9.9327 | -104.7868
9.9364 | -104.7614
9.9401 | -104.7375
9.9438 | -104.712
9.9474 | -104.6875
9.9511 | -104.6675
9.9548 | -104.6464
9.9585 | -104.6279
9.9622 | -104.6122
9.9659 | -104.5996
9.9696 | -104.5898
9.9733 | -104.5864
9.9769 | -104.5893
9.9806 | -104.5989
9.9843 | -104.6111
9.988 | -104.6249
9.9917 | -104.6422
9.9954 | -104.665
9.9991 | -104.6896
10.0028 | -104.715
10.0065 | -104.7372
10.0101 | -104.761
10.0138 | -104.787
10.0175 | -104.8132
10.0212 | -104.8369
10.0249 | -104.8589
10.0286 | -104.8793
10.0323 | -104.895
10.036 | -104.9137
10.0396 | -104.9301
10.0433 | -104.9422
10.047 | -104.9541
10.0507 | -104.9621
10.0544 | -104.9654
10.0581 | -104.9691
10.0618 | -104.9762
10.0655 | -104.9832
10.0692 | -104.9826
10.0728 | -104.9776
10.0765 | -104.9762
10.0802 | -104.9782
10.0839 | -104.9778
10.0876 | -104.9744
10.0913 | -104.9684
10.095 | -104.9586
10.0987 | -104.9468
10.1024 | -104.9313
10.106 | -104.9136
10.1097 | -104.8884
10.1134 | -104.8699
10.1171 | -104.8503
10.1208 | -104.8277
10.1245 | -104.8073
10.1282 | -104.7863
10.1319 | -104.7663
10.1355 | -104.7512
10.1392 | -104.7408
10.1429 | -104.7307
10.1466 | -104.7246
10.1503 | -104.7209
10.154 | -104.7225
10.1577 | -104.7244
10.1614 | -104.7273
10.1651 | -104.7349
10.1687 | -104.7393
10.1724 | -104.7456
10.1761 | -104.755
10.1798 | -104.7674
10.1835 | -104.7764
10.1872 | -104.7848
10.1909 | -104.7927
10.1946 | -104.8024
10.1982 | -104.8129
10.2019 | -104.8252
10.2056 | -104.836
10.2093 | -104.8402
10.213 | -104.8516
10.2167 | -104.8646
10.2204 | -104.8787
10.2241 | -104.8943
10.2278 | -104.9048
10.2314 | -104.9137
10.2351 | -104.9259
10.2388 | -104.9376
10.2425 | -104.9459
10.2462 | -104.9585
10.2499 | -104.9739
10.2536 | -104.9883
10.2573 | -105.0027
10.2609 | -105.015
10.2646 | -105.028
10.2683 | -105.0414
10.272 | -105.0582
10.2757 | -105.0724
10.2794 | -105.0866
10.2831 | -105.1037
10.2868 | -105.1147
10.2905 | -105.1159
10.2941 | -105.1207
10.2978 | -105.1246
10.3015 | -105.1218
10.3052 | -105.1142
10.3089 | -105.1061
10.3126 | -105.0961
10.3163 | -105.0782
10.32 | -105.0618
10.3237 | -105.0447
10.3273 | -105.0229
10.331 | -105.0012
10.3347 | -104.9807
10.3384 | -104.9582
10.3421 | -104.9388
10.3458 | -104.9179
10.3495 | -104.9003
10.3532 | -104.8846
10.3568 | -104.8724
10.3605 | -104.8652
10.3642 | -104.8554
10.3679 | -104.8464
10.3716 | -104.8403
10.3753 | -104.8367
10.379 | -104.8357
10.3827 | -104.835
10.3864 | -104.8333
10.39 | -104.8339
10.3937 | -104.8337
10.3974 | -104.8318
10.4011 | -104.8288
10.4048 | -104.8282
10.4085 | -104.8306
10.4122 | -104.8332
10.4159 | -104.8363
10.4195 | -104.8388
10.4232 | -104.8419
10.4269 | -104.8498
10.4306 | -104.86
10.4343 | -104.8692
10.438 | -104.8823
10.4417 | -104.901
10.4454 | -104.918
10.4491 | -104.9364
10.4527 | -104.956
10.4564 | -104.9747
10.4601 | -104.9925
10.4638 | -105.0118
10.4675 | -105.0329
10.4712 | -105.0539
10.4749 | -105.0701
10.4786 | -105.0901
10.4823 | -105.1082
10.4859 | -105.1251
10.4896 | -105.1452
10.4933 | -105.1592
10.497 | -105.1702
10.5007 | -105.1767
10.5044 | -105.1793
10.5081 | -105.1839
10.5118 | -105.1866
10.5154 | -105.185
10.5191 | -105.1829
10.5228 | -105.1799
10.5265 | -105.1739
10.5302 | -105.1684
10.5339 | -105.1625
10.5376 | -105.1539
10.5413 | -105.1465
10.545 | -105.1427
10.5486 | -105.1358
10.5523 | -105.1256
10.556 | -105.1138
10.5597 | -105.105
10.5634 | -105.0962
10.5671 | -105.091
10.5708 | -105.0867
10.5745 | -105.0806
10.5781 | -105.0732
10.5818 | -105.0675
10.5855 | -105.0643
10.5892 | -105.0612
10.5929 | -105.0584
10.5966 | -105.0527
10.6003 | -105.0486
10.604 | -105.0467
10.6077 | -105.0447
10.6113 | -105.0379
10.615 | -105.0254
10.6187 | -105.0144
10.6224 | -104.9982
10.6261 | -104.9819
10.6298 | -104.9648
10.6335 | -104.9479
10.6372 | -104.9321
10.6408 | -104.9178
10.6445 | -104.9007
10.6482 | -104.8818
10.6519 | -104.8666
10.6556 | -104.8528
10.6593 | -104.8427
10.663 | -104.8305
10.6667 | -104.8174
10.6704 | -104.8097
10.674 | -104.8036
10.6777 | -104.792
10.6814 | -104.7827
10.6851 | -104.7752
10.6888 | -104.7666
10.6925 | -104.7607
10.6962 | -104.7546
10.6999 | -104.7486
10.7035 | -104.7433
10.7072 | -104.7429
10.7109 | -104.7439
10.7146 | -104.7453
10.7183 | -104.7508
10.722 | -104.7574
10.7257 | -104.769
10.7294 | -104.7811
10.7331 | -104.7928
10.7367 | -104.805
10.7404 | -104.8188
10.7441 | -104.8353
10.7478 | -104.8501
10.7515 | -104.8673
10.7552 | -104.8848
10.7589 | -104.9024
10.7626 | -104.9239
10.7663 | -104.9475
10.7699 | -104.9685
10.7736 | -104.9844
10.7773 | -105.0038
10.781 | -105.022
10.7847 | -105.0364
10.7884 | -105.0491
10.7921 | -105.0605
10.7958 | -105.072
10.7994 | -105.0802
10.8031 | -105.0906
10.8068 | -105.1011
10.8105 | -105.1073
10.8142 | -105.1141
10.8179 | -105.1206
10.8216 | -105.127
10.8253 | -105.1393
10.829 | -105.1494
10.8326 | -105.159
10.8363 | -105.1751
10.84 | -105.1902
10.8437 | -105.2016
10.8474 | -105.2155
10.8511 | -105.2295
10.8548 | -105.2417
10.8585 | -105.2537
10.8621 | -105.2666
10.8658 | -105.2738
10.8695 | -105.2787
10.8732 | -105.2812
10.8769 | -105.2821
10.8806 | -105.278
10.8843 | -105.2732
10.888 | -105.2712
10.8917 | -105.2685
10.8953 | -105.2625
10.899 | -105.2534
10.9027 | -105.2427
10.9064 | -105.2332
10.9101 | -105.2182
10.9138 | -105.2057
10.9175 | -105.1963
10.9212 | -105.1875
10.9248 | -105.1799
10.9285 | -105.1781
10.9322 | -105.1783
10.9359 | -105.179
10.9396 | -105.1837
10.9433 | -105.1893
10.947 | -105.1991
10.9507 | -105.2115
10.9544 | -105.2245
10.958 | -105.2395
10.9617 | -105.2571
10.9654 | -105.2767
10.9691 | -105.293
10.9728 | -105.3079
10.9765 | -105.3231
10.9802 | -105.338
10.9839 | -105.3472
10.9876 | -105.3575
10.9912 | -105.3638
10.9949 | -105.3666
10.9986 | -105.3649
11.0023 | -105.3628
11.006 | -105.3551
11.0097 | -105.341
11.0134 | -105.3213
11.0171 | -105.3017
11.0207 | -105.2775
11.0244 | -105.2517
11.0281 | -105.2255
11.0318 | -105.2014
11.0355 | -105.1751
11.0392 | -105.1493
11.0429 | -105.122
11.0466 | -105.1004
11.0503 | -105.0806
11.0539 | -105.0573
11.0576 | -105.0371
11.0613 | -105.0183
11.065 | -104.9999
11.0687 | -104.9855
11.0724 | -104.9696
11.0761 | -104.9532
11.0798 | -104.937
11.0834 | -104.9242
11.0871 | -104.9107
11.0908 | -104.8976
11.0945 | -104.8849
11.0982 | -104.873
11.1019 | -104.8636
11.1056 | -104.854
11.1093 | -104.8465
11.113 | -104.8379
11.1166 | -104.8287
11.1203 | -104.8181
11.124 | -104.8117
11.1277 | -104.8125
11.1314 | -104.8145
11.1351 | -104.8165
11.1388 | -104.8203
11.1425 | -104.8276
11.1462 | -104.8373
11.1498 | -104.8517
11.1535 | -104.8626
11.1572 | -104.874
11.1609 | -104.8895
11.1646 | -104.9083
11.1683 | -104.9289
11.172 | -104.9508
11.1757 | -104.9728
11.1793 | -104.9918
11.183 | -105.0118
11.1867 | -105.0293
11.1904 | -105.0439
11.1941 | -105.0554
11.1978 | -105.0721
11.2015 | -105.0902
11.2052 | -105.1058
11.2089 | -105.1142
11.2125 | -105.1238
11.2162 | -105.1375
11.2199 | -105.1481
11.2236 | -105.153
11.2273 | -105.1531
11.231 | -105.1528
11.2347 | -105.1505
11.2384 | -105.1477
11.242 | -105.1441
11.2457 | -105.1408
11.2494 | -105.1366
11.2531 | -105.1313
11.2568 | -105.1251
11.2605 | -105.1156
11.2642 | -105.1092
11.2679 | -105.1031
11.2716 | -105.0963
11.2752 | -105.0946
11.2789 | -105.0975
11.2826 | -105.0997
11.2863 | -105.0983
11.29 | -105.099
11.2937 | -105.104
11.2974 | -105.1077
11.3011 | -105.1123
11.3047 | -105.1184
11.3084 | -105.1239
11.3121 | -105.1298
11.3158 | -105.1392
11.3195 | -105.1481
11.3232 | -105.151
11.3269 | -105.1543
11.3306 | -105.1585
11.3343 | -105.1627
11.3379 | -105.1677
11.3416 | -105.1735
11.3453 | -105.181
11.349 | -105.1871
11.3527 | -105.1941
11.3564 | -105.2026
11.3601 | -105.2072
11.3638 | -105.2066
11.3674 | -105.2094
11.3711 | -105.2152
11.3748 | -105.2232
11.3785 | -105.2276
11.3822 | -105.2323
11.3859 | -105.2371
11.3896 | -105.2468
11.3933 | -105.2576
11.397 | -105.2634
11.4006 | -105.2656
11.4043 | -105.2661
11.408 | -105.2636
11.4117 | -105.2589
11.4154 | -105.2548
11.4191 | -105.251
11.4228 | -105.2454
11.4265 | -105.2375
11.4302 | -105.223
11.4338 | -105.2095
11.4375 | -105.196
11.4412 | -105.1771
11.4449 | -105.1541
11.4486 | -105.135
11.4523 | -105.1133
11.456 | -105.0874
11.4597 | -105.0617
11.4633 | -105.0413
11.467 | -105.026
11.4707 | -105.0148
11.4744 | -105.0031
11.4781 | -104.9906
11.4818 | -104.9812
11.4855 | -104.9726
11.4892 | -104.9641
11.4929 | -104.9563
11.4965 | -104.9476
11.5002 | -104.939
11.5039 | -104.9311
11.5076 | -104.9216
11.5113 | -104.9141
11.515 | -104.9091
11.5187 | -104.9059
11.5224 | -104.902
11.526 | -104.8995
11.5297 | -104.8999
11.5334 | -104.8991
11.5371 | -104.8978
11.5408 | -104.9009
11.5445 | -104.9097
11.5482 | -104.9217
11.5519 | -104.935
11.5556 | -104.9485
11.5592 | -104.9626
11.5629 | -104.9819
11.5666 | -105.0053
11.5703 | -105.0294
11.574 | -105.0566
11.5777 | -105.084
11.5814 | -105.1122
11.5851 | -105.1429
11.5888 | -105.1726
11.5924 | -105.203
11.5961 | -105.236
11.5998 | -105.269
11.6035 | -105.3021
11.6072 | -105.3361
11.6109 | -105.368
11.6146 | -105.3982
11.6183 | -105.4291
11.6219 | -105.4563
11.6256 | -105.4785
11.6293 | -105.4938
11.633 | -105.4981
11.6367 | -105.4949
11.6404 | -105.489
11.6441 | -105.4796
11.6478 | -105.4637
11.6515 | -105.4405
11.6551 | -105.4135
11.6588 | -105.382
11.6625 | -105.3465
11.6662 | -105.3086
11.6699 | -105.2699
11.6736 | -105.233
11.6773 | -105.1966
11.681 | -105.1613
11.6846 | -105.1297
11.6883 | -105.1028
11.692 | -105.0789
11.6957 | -105.0573
11.6994 | -105.041
11.7031 | -105.0489
11.7068 | -105.0424
11.7105 | -105.0348
11.7142 | -105.0304
11.7178 | -105.0307
11.7215 | -105.0336
11.7252 | -105.0372
11.7289 | -105.0459
11.7326 | -105.0558
11.7363 | -105.0657
11.74 | -105.0768
11.7437 | -105.0878
11.7473 | -105.0975
11.751 | -105.1062
11.7547 | -105.1144
11.7584 | -105.1248
11.7621 | -105.1369
11.7658 | -105.1466
11.7695 | -105.157
11.7732 | -105.1671
11.7769 | -105.1777
11.7805 | -105.1912
11.7842 | -105.2034
11.7879 | -105.2157
11.7916 | -105.2329
11.7953 | -105.249
11.799 | -105.2661
11.8027 | -105.2869
11.8064 | -105.3065
11.8101 | -105.325
11.8137 | -105.3447
11.8174 | -105.3647
11.8211 | -105.3879
11.8248 | -105.4098
11.8285 | -105.4295
11.8322 | -105.4492
11.8359 | -105.4687
11.8396 | -105.484
11.8432 | -105.4944
11.8469 | -105.5022
11.8506 | -105.5093
11.8543 | -105.5146
11.858 | -105.5149
11.8617 | -105.5146
11.8654 | -105.5097
11.8691 | -105.5014
11.8728 | -105.4881
11.8764 | -105.4736
11.8801 | -105.4574
11.8838 | -105.4422
11.8875 | -105.4244
11.8912 | -105.401
11.8949 | -105.3772
11.8986 | -105.3561
11.9023 | -105.3331
11.9059 | -105.3123
11.9096 | -105.2937
11.9133 | -105.2772
11.917 | -105.2614
11.9207 | -105.2463
11.9244 | -105.2328
11.9281 | -105.2198
11.9318 | -105.21
11.9355 | -105.2011
11.9391 | -105.1956
11.9428 | -105.1891
11.9465 | -105.1816
11.9502 | -105.1785
11.9539 | -105.1726
11.9576 | -105.1652
11.9613 | -105.1626
11.965 | -105.1595
11.9686 | -105.1606
11.9723 | -105.1647
11.976 | -105.171
11.9797 | -105.1794
11.9834 | -105.1906
11.9871 | -105.2008
11.9908 | -105.2124
11.9945 | -105.2241
11.9982 | -105.2397
12.0018 | -105.2545
12.0055 | -105.2707
12.0092 | -105.2859
12.0129 | -105.3038
12.0166 | -105.3269
12.0203 | -105.3487
12.024 | -105.3671
12.0277 | -105.3838
12.0314 | -105.4023
12.035 | -105.4202
12.0387 | -105.436
12.0424 | -105.4489
12.0461 | -105.4603
12.0498 | -105.4704
12.0535 | -105.4795
12.0572 | -105.4872
12.0609 | -105.4904
12.0645 | -105.4929
12.0682 | -105.4968
12.0719 | -105.5012
12.0756 | -105.5037
12.0793 | -105.5004
12.083 | -105.496
12.0867 | -105.4908
12.0904 | -105.4859
12.0941 | -105.4824
12.0977 | -105.4799
12.1014 | -105.4758
12.1051 | -105.4691
12.1088 | -105.4658
12.1125 | -105.4641
12.1162 | -105.4615
12.1199 | -105.4618
12.1236 | -105.46
12.1272 | -105.4595
12.1309 | -105.4618
12.1346 | -105.4626
12.1383 | -105.462
12.142 | -105.4598
12.1457 | -105.4579
12.1494 | -105.4527
12.1531 | -105.45
12.1568 | -105.4477
12.1604 | -105.4465
12.1641 | -105.4422
12.1678 | -105.4396
12.1715 | -105.4355
12.1752 | -105.4301
12.1789 | -105.4232
12.1826 | -105.4158
12.1863 | -105.404
12.1899 | -105.3948
12.1936 | -105.3835
12.1973 | -105.3722
12.201 | -105.3653
12.2047 | -105.3577
12.2084 | -105.3485
12.2121 | -105.3411
12.2158 | -105.3335
12.2195 | -105.3266
12.2231 | -105.3215
12.2268 | -105.316
12.2305 | -105.3085
12.2342 | -105.3043
12.2379 | -105.3105
12.2416 | -105.3112
12.2453 | -105.3099
12.249 | -105.3071
12.2527 | -105.3072
12.2563 | -105.3106
12.26 | -105.3179
12.2637 | -105.3231
12.2674 | -105.3268
12.2711 | -105.3305
12.2748 | -105.3335
12.2785 | -105.3364
12.2822 | -105.3404
12.2858 | -105.3396
12.2895 | -105.3373
12.2932 | -105.338
12.2969 | -105.338
12.3006 | -105.3374
12.3043 | -105.3365
12.308 | -105.3327
12.3117 | -105.3284
12.3154 | -105.3244
12.319 | -105.3191
12.3227 | -105.3152
12.3264 | -105.3153
12.3301 | -105.3139
12.3338 | -105.3131
12.3375 | -105.3118
12.3412 | -105.3112
12.3449 | -105.3095
12.3485 | -105.3048
12.3522 | -105.3032
12.3559 | -105.3016
12.3596 | -105.2993
12.3633 | -105.2993
12.367 | -105.2931
12.3707 | -105.2908
12.3744 | -105.2898
12.3781 | -105.2838
12.3817 | -105.279
12.3854 | -105.2767
12.3891 | -105.2761
12.3928 | -105.2767
12.3965 | -105.2751
12.4002 | -105.2744
12.4039 | -105.2741
12.4076 | -105.2755
12.4112 | -105.2789
12.4149 | -105.2775
12.4186 | -105.2735
12.4223 | -105.2711
12.426 | -105.2709
12.4297 | -105.2723
12.4334 | -105.2721
12.4371 | -105.2722
12.4408 | -105.2699
12.4444 | -105.2724
12.4481 | -105.28
12.4518 | -105.2883
12.4555 | -105.2953
12.4592 | -105.3042
12.4629 | -105.3137
12.4666 | -105.3238
12.4703 | -105.3347
12.474 | -105.3437
12.4776 | -105.353
12.4813 | -105.3641
12.485 | -105.3765
12.4887 | -105.3899
12.4924 | -105.4017
12.4961 | -105.4111
12.4998 | -105.4236
12.5035 | -105.4336
12.5071 | -105.4424
12.5108 | -105.4515
12.5145 | -105.4594
12.5182 | -105.4687
12.5219 | -105.4783
12.5256 | -105.4853
12.5293 | -105.4926
12.533 | -105.4977
12.5367 | -105.5015
12.5403 | -105.5038
12.544 | -105.5052
12.5477 | -105.5071
12.5514 | -105.506
12.5551 | -105.5063
12.5588 | -105.5038
12.5625 | -105.5022
12.5662 | -105.5015
12.5698 | -105.5017
12.5735 | -105.4998
12.5772 | -105.4995
12.5809 | -105.4982
12.5846 | -105.4985
12.5883 | -105.4995
12.592 | -105.4996
12.5957 | -105.5017
12.5994 | -105.5081
12.603 | -105.514
12.6067 | -105.5143
12.6104 | -105.5146
12.6141 | -105.5165
12.6178 | -105.5186
12.6215 | -105.5211
12.6252 | -105.5235
12.6289 | -105.5234
12.6326 | -105.5219
12.6362 | -105.5233
12.6399 | -105.5264
12.6436 | -105.5296
12.6473 | -105.5354
12.651 | -105.5426
12.6547 | -105.5483
12.6584 | -105.554
12.6621 | -105.5595
12.6657 | -105.5662
12.6694 | -105.5744
12.6731 | -105.5821
12.6768 | -105.5929
12.6805 | -105.6036
12.6842 | -105.6147
12.6879 | -105.6265
12.6916 | -105.6384
12.6953 | -105.6528
12.6989 | -105.6701
12.7026 | -105.684
12.7063 | -105.6991
12.71 | -105.7124
12.7137 | -105.7231
12.7174 | -105.7367
12.7211 | -105.749
12.7248 | -105.7594
12.7284 | -105.7673
12.7321 | -105.7741
12.7358 | -105.7739
12.7395 | -105.7746
12.7432 | -105.7743
12.7469 | -105.771
12.7506 | -105.7634
12.7543 | -105.7556
12.758 | -105.7483
12.7616 | -105.7373
12.7653 | -105.7256
12.769 | -105.6999
12.7727 | -105.6833
12.7764 | -105.6662
12.7801 | -105.6495
12.7838 | -105.6308
12.7875 | -105.6111
12.7911 | -105.5942
12.7948 | -105.5762
12.7985 | -105.5569
12.8022 | -105.5356
12.8059 | -105.5174
12.8096 | -105.499
12.8133 | -105.483
12.817 | -105.47
12.8207 | -105.4564
12.8243 | -105.4461
12.828 | -105.4395
12.8317 | -105.4319
12.8354 | -105.4223
12.8391 | -105.4147
12.8428 | -105.4104
12.8465 | -105.4068
12.8502 | -105.4002
12.8538 | -105.3937
12.8575 | -105.387
12.8612 | -105.3795
12.8649 | -105.3678
12.8686 | -105.3571
12.8723 | -105.3451
12.876 | -105.3294
12.8797 | -105.3115
12.8834 | -105.2914
12.887 | -105.2743
12.8907 | -105.2549
12.8944 | -105.2362
12.8981 | -105.2163
12.9018 | -105.1935
12.9055 | -105.1704
12.9092 | -105.1478
12.9129 | -105.1288
12.9166 | -105.1106
12.9202 | -105.0964
12.9239 | -105.0825
12.9276 | -105.069
12.9313 | -105.0599
12.935 | -105.0537
12.9387 | -105.0494
12.9424 | -105.0482
12.9461 | -105.0484
12.9497 | -105.0493
12.9534 | -105.0559
12.9571 | -105.0672
12.9608 | -105.076
12.9645 | -105.085
12.9682 | -105.0986
12.9719 | -105.1153
12.9756 | -105.1335
12.9793 | -105.1499
12.9829 | -105.1668
12.9866 | -105.1845
12.9903 | -105.2039
12.994 | -105.2214
12.9977 | -105.2396
13.0014 | -105.2557
13.0051 | -105.2707
13.0088 | -105.2856
13.0124 | -105.2982
13.0161 | -105.3073
13.0198 | -105.3189
13.0235 | -105.3314
13.0272 | -105.343
13.0309 | -105.3477
13.0346 | -105.3526
13.0383 | -105.3611
13.042 | -105.3683
13.0456 | -105.3767
13.0493 | -105.384
13.053 | -105.3897
13.0567 | -105.3938
13.0604 | -105.3999
13.0641 | -105.4078
13.0678 | -105.413
13.0715 | -105.416
13.0752 | -105.4189
13.0788 | -105.4216
13.0825 | -105.4252
13.0862 | -105.4282
13.0899 | -105.4286
13.0936 | -105.4298
13.0973 | -105.4284
13.101 | -105.4257
13.1047 | -105.4243
13.1083 | -105.4209
13.112 | -105.4173
13.1157 | -105.4102
13.1194 | -105.4052
13.1231 | -105.4016
13.1268 | -105.3987
13.1305 | -105.3935
13.1342 | -105.3921
13.1379 | -105.3949
13.1415 | -105.3958
13.1452 | -105.3926
13.1489 | -105.3933
13.1526 | -105.3958
13.1563 | -105.3979
13.16 | -105.398
13.1637 | -105.4023
13.1674 | -105.4094
13.171 | -105.4164
13.1747 | -105.4212
13.1784 | -105.4271
13.1821 | -105.4359
13.1858 | -105.442
13.1895 | -105.4489
13.1932 | -105.4547
13.1969 | -105.4584
13.2006 | -105.4642
13.2042 | -105.4708
13.2079 | -105.4728
13.2116 | -105.4745
13.2153 | -105.4754
13.219 | -105.4755
13.2227 | -105.4703
13.2264 | -105.4633
13.2301 | -105.4533
13.2337 | -105.4411
13.2374 | -105.4273
13.2411 | -105.411
13.2448 | -105.3924
13.2485 | -105.3703
13.2522 | -105.3464
13.2559 | -105.3249
13.2596 | -105.3031
13.2633 | -105.2791
13.2669 | -105.256
13.2706 | -105.2328
13.2743 | -105.2107
13.278 | -105.1916
13.2817 | -105.1717
13.2854 | -105.1533
13.2891 | -105.1328
13.2928 | -105.1132
13.2965 | -105.0964
13.3001 | -105.0868
13.3038 | -105.0736
13.3075 | -105.0629
13.3112 | -105.0483
13.3149 | -105.038
13.3186 | -105.0293
13.3223 | -105.0202
13.326 | -105.0102
13.3296 | -104.9975
13.3333 | -104.9861
13.337 | -104.9788
13.3407 | -104.9677
13.3444 | -104.9555
13.3481 | -104.9427
13.3518 | -104.9328
13.3555 | -104.921
13.3592 | -104.9089
13.3628 | -104.8969
13.3665 | -104.8861
13.3702 | -104.8757
13.3739 | -104.8671
13.3776 | -104.8619
13.3813 | -104.8588
13.385 | -104.8555
13.3887 | -104.8537
13.3923 | -104.8535
13.396 | -104.8526
13.3997 | -104.855
13.4034 | -104.861
13.4071 | -104.8675
13.4108 | -104.8744
13.4145 | -104.8836
13.4182 | -104.8946
13.4219 | -104.9089
13.4255 | -104.9207
13.4292 | -104.9345
13.4329 | -104.9529
13.4366 | -104.9713
13.4403 | -104.9895
13.444 | -105.0064
13.4477 | -105.0227
13.4514 | -105.0408
13.455 | -105.0621
13.4587 | -105.0834
13.4624 | -105.1027
13.4661 | -105.1213
13.4698 | -105.1395
13.4735 | -105.159
13.4772 | -105.1787
13.4809 | -105.2011
13.4846 | -105.2225
13.4882 | -105.2407
13.4919 | -105.2539
13.4956 | -105.2687
13.4993 | -105.2822
13.503 | -105.2936
13.5067 | -105.3068
13.5104 | -105.3184
13.5141 | -105.3295
13.5177 | -105.3428
13.5214 | -105.3555
13.5251 | -105.3653
13.5288 | -105.3757
13.5325 | -105.3867
13.5362 | -105.4004
13.5399 | -105.4122
13.5436 | -105.4227
13.5473 | -105.4334
13.5509 | -105.4444
13.5546 | -105.456
13.5583 | -105.467
13.562 | -105.4775
13.5657 | -105.4863
13.5694 | -105.4918
13.5731 | -105.498
13.5768 | -105.5043
13.5805 | -105.5124
13.5841 | -105.5198
13.5878 | -105.5239
13.5915 | -105.5306
13.5952 | -105.5368
13.5989 | -105.5384
13.6026 | -105.5433
13.6063 | -105.5498
13.61 | -105.5538
13.6136 | -105.5564
13.6173 | -105.5576
13.621 | -105.5638
13.6247 | -105.567
13.6284 | -105.57
13.6321 | -105.5741
13.6358 | -105.5762
13.6395 | -105.5782
13.6432 | -105.5778
13.6468 | -105.5756
13.6505 | -105.5713
13.6542 | -105.5661
13.6579 | -105.5618
13.6616 | -105.5572
13.6653 | -105.5494
13.669 | -105.5421
13.6727 | -105.5336
13.6763 | -105.5245
13.68 | -105.5138
13.6837 | -105.5052
13.6874 | -105.4982
13.6911 | -105.4914
13.6948 | -105.4837
13.6985 | -105.4773
13.7022 | -105.4718
13.7059 | -105.467
13.7095 | -105.4673
13.7132 | -105.4682
13.7169 | -105.4685
13.7206 | -105.4747
13.7243 | -105.4821
13.728 | -105.4866
13.7317 | -105.4898
13.7354 | -105.4953
13.7391 | -105.5019
13.7427 | -105.5091
13.7464 | -105.519
13.7501 | -105.5291
13.7538 | -105.5405
13.7575 | -105.551
13.7612 | -105.5602
13.7649 | -105.5708
13.7686 | -105.579
13.7722 | -105.5849
13.7759 | -105.5926
13.7796 | -105.6033
13.7833 | -105.6131
13.787 | -105.6246
13.7907 | -105.6344
13.7944 | -105.6466
13.7981 | -105.6568
13.8018 | -105.6644
13.8054 | -105.6695
13.8091 | -105.6719
13.8128 | -105.674
13.8165 | -105.6746
13.8202 | -105.6735
13.8239 | -105.665
13.8276 | -105.6526
13.8313 | -105.6291
13.8349 | -105.6173
13.8386 | -105.6048
13.8423 | -105.5919
13.846 | -105.5739
13.8497 | -105.551
13.8534 | -105.529
13.8571 | -105.5048
13.8608 | -105.4805
13.8645 | -105.4585
13.8681 | -105.4368
13.8718 | -105.4151
13.8755 | -105.3928
13.8792 | -105.3745
13.8829 | -105.357
13.8866 | -105.3409
13.8903 | -105.3261
13.894 | -105.3141
13.8976 | -105.3063
13.9013 | -105.2994
13.905 | -105.2942
13.9087 | -105.2895
13.9124 | -105.2883
13.9161 | -105.292
13.9198 | -105.2955
13.9235 | -105.2989
13.9272 | -105.3067
13.9308 | -105.3178
13.9345 | -105.3271
13.9382 | -105.3346
13.9419 | -105.3411
13.9456 | -105.3497
13.9493 | -105.3588
13.953 | -105.3696
13.9567 | -105.3751
13.9604 | -105.3568
13.964 | -105.3305
13.9677 | -105.3194
13.9714 | -105.311
13.9751 | -105.296
13.9788 | -105.2757
13.9825 | -105.2503
13.9862 | -105.2192
13.9899 | -105.1833
13.9935 | -105.1416
13.9972 | -105.0987
14.0009 | -105.0555
14.0046 | -105.0083
14.0083 | -104.9548
14.012 | -104.8964
14.0157 | -104.838
14.0194 | -104.7784
14.0231 | -104.7211
14.0267 | -104.6637
14.0304 | -104.6064
14.0341 | -104.5482
14.0378 | -104.495
14.0415 | -104.4455
14.0452 | -104.4012
14.0489 | -104.3621
14.0526 | -104.3326
14.0562 | -104.309
14.0599 | -104.2927
14.0636 | -104.2833
14.0673 | -104.2792
14.071 | -104.2834
14.0747 | -104.2944
14.0784 | -104.3127
14.0821 | -104.3387
14.0858 | -104.3708
14.0894 | -104.4092
14.0931 | -104.4522
14.0968 | -104.4993
14.1005 | -104.5497
14.1042 | -104.6041
14.1079 | -104.6612
14.1116 | -104.7197
14.1153 | -104.7799
14.1189 | -104.8398
14.1226 | -104.8988
14.1263 | -104.9587
14.13 | -105.0168
14.1337 | -105.0702
14.1374 | -105.1271
14.1411 | -105.1815
14.1448 | -105.2338
14.1485 | -105.2808
14.1521 | -105.3265
14.1558 | -105.3682
14.1595 | -105.415
14.1632 | -105.4769
14.1669 | -105.5308
14.1706 | -105.5618
14.1743 | -105.5857
14.178 | -105.6068
14.1817 | -105.6268
14.1853 | -105.6467
14.189 | -105.6652
14.1927 | -105.6828
14.1964 | -105.6932
14.2001 | -105.6994
14.2038 | -105.704
14.2075 | -105.7113
14.2112 | -105.7178
14.2148 | -105.7229
14.2185 | -105.724
14.2222 | -105.7239
14.2259 | -105.7234
14.2296 | -105.7219
14.2333 | -105.7216
14.237 | -105.723
14.2407 | -105.7206
14.2444 | -105.7188
14.248 | -105.7165
14.2517 | -105.7177
14.2554 | -105.7157
14.2591 | -105.7148
14.2628 | -105.715
14.2665 | -105.7162
14.2702 | -105.7186
14.2739 | -105.7205
14.2775 | -105.7261
14.2812 | -105.7314
14.2849 | -105.7342
14.2886 | -105.7402
14.2923 | -105.748
14.296 | -105.7575
14.2997 | -105.7686
14.3034 | -105.777
14.3071 | -105.7835
14.3107 | -105.7916
14.3144 | -105.8005
14.3181 | -105.8072
14.3218 | -105.8146
14.3255 | -105.82
14.3292 | -105.8259
14.3329 | -105.829
14.3366 | -105.8312
14.3402 | -105.8298
14.3439 | -105.8274
14.3476 | -105.8238
14.3513 | -105.821
14.355 | -105.8164
14.3587 | -105.8104
14.3624 | -105.8017
14.3661 | -105.7921
14.3698 | -105.7827
14.3734 | -105.7779
14.3771 | -105.7716
14.3808 | -105.7633
14.3845 | -105.7559
14.3882 | -105.746
14.3919 | -105.7384
14.3956 | -105.7311
14.3993 | -105.7216
14.403 | -105.7169
14.4066 | -105.7124
14.4103 | -105.7055
14.414 | -105.6963
14.4177 | -105.6882
14.4214 | -105.6764
14.4251 | -105.6639
14.4288 | -105.6521
14.4325 | -105.6395
14.4361 | -105.6261
14.4398 | -105.6114
14.4435 | -105.596
14.4472 | -105.581
14.4509 | -105.5665
14.4546 | -105.5535
14.4583 | -105.539
14.462 | -105.5225
14.4657 | -105.5068
14.4693 | -105.4918
14.473 | -105.4777
14.4767 | -105.4653
14.4804 | -105.4536
14.4841 | -105.443
14.4878 | -105.4366
14.4915 | -105.4316
14.4952 | -105.4255
14.4988 | -105.4181
14.5025 | -105.4142
14.5062 | -105.4119
14.5099 | -105.4139
14.5136 | -105.4181
14.5173 | -105.4194
14.521 | -105.4247
14.5247 | -105.4333
14.5284 | -105.4414
14.532 | -105.4509
14.5357 | -105.459
14.5394 | -105.4708
14.5431 | -105.4847
14.5468 | -105.4949
14.5505 | -105.5031
14.5542 | -105.5124
14.5579 | -105.5173
14.5615 | -105.5219
14.5652 | -105.53
14.5689 | -105.5366
14.5726 | -105.5383
14.5763 | -105.5429
14.58 | -105.547
14.5837 | -105.5465
14.5874 | -105.5452
14.5911 | -105.5464
14.5947 | -105.5459
14.5984 | -105.5466
14.6021 | -105.5435
14.6058 | -105.5413
14.6095 | -105.5439
14.6132 | -105.5427
14.6169 | -105.5388
14.6206 | -105.5388
14.6243 | -105.5404
14.6279 | -105.5434
14.6316 | -105.5501
14.6353 | -105.5564
14.639 | -105.5641
14.6427 | -105.5703
14.6464 | -105.5787
14.6501 | -105.5887
14.6538 | -105.5996
14.6574 | -105.6082
14.6611 | -105.6148
14.6648 | -105.6241
14.6685 | -105.636
14.6722 | -105.6504
14.6759 | -105.6614
14.6796 | -105.6723
14.6833 | -105.6842
14.687 | -105.6947
14.6906 | -105.7038
14.6943 | -105.7122
14.698 | -105.721
14.7017 | -105.7287
14.7054 | -105.7376
14.7091 | -105.7464
14.7128 | -105.7539
14.7165 | -105.759
14.7201 | -105.7617
14.7238 | -105.7651
14.7275 | -105.7697
14.7312 | -105.7694
14.7349 | -105.7649
14.7386 | -105.7627
14.7423 | -105.7606
14.746 | -105.7574
14.7497 | -105.7514
14.7533 | -105.7463
14.757 | -105.7436
14.7607 | -105.7417
14.7644 | -105.736
14.7681 | -105.7346
14.7718 | -105.7314
14.7755 | -105.7229
14.7792 | -105.7174
14.7828 | -105.7146
14.7865 | -105.7112
14.7902 | -105.7073
14.7939 | -105.7058
14.7976 | -105.7026
14.8013 | -105.6976
14.805 | -105.6948
14.8087 | -105.691
14.8124 | -105.6885
14.816 | -105.6848
14.8197 | -105.6793
14.8234 | -105.6793
14.8271 | -105.6814
14.8308 | -105.68
14.8345 | -105.6773
14.8382 | -105.6736
14.8419 | -105.6718
14.8456 | -105.6694
14.8492 | -105.6646
14.8529 | -105.66
14.8566 | -105.6541
14.8603 | -105.647
14.864 | -105.6395
14.8677 | -105.6308
14.8714 | -105.6166
14.8751 | -105.6008
14.8787 | -105.5851
14.8824 | -105.5687
14.8861 | -105.5533
14.8898 | -105.5355
14.8935 | -105.5157
14.8972 | -105.4888
14.9009 | -105.4692
14.9046 | -105.4495
14.9083 | -105.4312
14.9119 | -105.4097
14.9156 | -105.3895
14.9193 | -105.37
14.923 | -105.3543
14.9267 | -105.3384
14.9304 | -105.3236
14.9341 | -105.3076
14.9378 | -105.2939
14.9414 | -105.2815
14.9451 | -105.2723
14.9488 | -105.2644
14.9525 | -105.256
14.9562 | -105.2493
14.9599 | -105.2446
14.9636 | -105.2409
14.9673 | -105.2362
14.971 | -105.2311
14.9746 | -105.2286
14.9783 | -105.2275
14.982 | -105.2232
14.9857 | -105.2139
14.9894 | -105.2059
14.9931 | -105.1978
14.9968 | -105.1894
15.0005 | -105.1803
15.0041 | -105.1701
15.0078 | -105.1591
15.0115 | -105.1518
15.0152 | -105.1437
15.0189 | -105.1334
15.0226 | -105.1209
15.0263 | -105.1033
15.03 | -105.0864
15.0337 | -105.0688
15.0373 | -105.0477
15.041 | -105.0257
15.0447 | -105.0047
15.0484 | -104.9835
15.0521 | -104.9642
15.0558 | -104.9411
15.0595 | -104.9202
15.0632 | -104.9029
15.0669 | -104.8876
15.0705 | -104.8729
15.0742 | -104.8572
15.0779 | -104.8444
15.0816 | -104.8354
15.0853 | -104.8315
15.089 | -104.825
15.0927 | -104.8224
15.0964 | -104.8243
15.1 | -104.8275
15.1037 | -104.8315
15.1074 | -104.8398
15.1111 | -104.8538
15.1148 | -104.8685
15.1185 | -104.8835
15.1222 | -104.8995
15.1259 | -104.9166
15.1296 | -104.9382
15.1332 | -104.9613
15.1369 | -104.9826
15.1406 | -105
15.1443 | -105.0185
15.148 | -105.0399
15.1517 | -105.0624
15.1554 | -105.0815
15.1591 | -105.0988
15.1627 | -105.115
15.1664 | -105.1328
15.1701 | -105.1456
15.1738 | -105.156
15.1775 | -105.1658
15.1812 | -105.1752
15.1849 | -105.1834
15.1886 | -105.1919
15.1923 | -105.1992
15.1959 | -105.2018
15.1996 | -105.2052
15.2033 | -105.2126
15.207 | -105.2169
15.2107 | -105.2169
15.2144 | -105.2147
15.2181 | -105.2121
15.2218 | -105.212
15.2255 | -105.2121
15.2291 | -105.2114
15.2328 | -105.2101
15.2365 | -105.2104
15.2402 | -105.2137
15.2439 | -105.2207
15.2476 | -105.2252
15.2513 | -105.229
15.255 | -105.2347
15.2586 | -105.2413
15.2623 | -105.2497
15.266 | -105.2605
15.2697 | -105.2713
15.2734 | -105.2857
15.2771 | -105.2985
15.2808 | -105.31
15.2845 | -105.3218
15.2882 | -105.3354
15.2918 | -105.351
15.2955 | -105.3666
15.2992 | -105.3833
15.3029 | -105.4012
15.3066 | -105.4173
15.3103 | -105.4304
15.314 | -105.442
15.3177 | -105.4546
15.3213 | -105.4693
15.325 | -105.4805
15.3287 | -105.4906
15.3324 | -105.5026
15.3361 | -105.5146
15.3398 | -105.5286
15.3435 | -105.5404
15.3472 | -105.5503
15.3509 | -105.5607
15.3545 | -105.5714
15.3582 | -105.5822
15.3619 | -105.5903
15.3656 | -105.5954
15.3693 | -105.6006
15.373 | -105.6044
15.3767 | -105.6082
15.3804 | -105.612
15.384 | -105.6149
15.3877 | -105.618
15.3914 | -105.6186
15.3951 | -105.6158
15.3988 | -105.6126
15.4025 | -105.6081
15.4062 | -105.6041
15.4099 | -105.5993
15.4136 | -105.5968
15.4172 | -105.5933
15.4209 | -105.5851
15.4246 | -105.5767
15.4283 | -105.5595
15.432 | -105.5524
15.4357 | -105.5413
15.4394 | -105.5298
15.4431 | -105.5135
15.4467 | -105.4948
15.4504 | -105.474
15.4541 | -105.4554
15.4578 | -105.438
15.4615 | -105.4171
15.4652 | -105.3949
15.4689 | -105.3705
15.4726 | -105.3453
15.4763 | -105.3214
15.4799 | -105.2956
15.4836 | -105.2687
15.4873 | -105.2412
15.491 | -105.2127
15.4947 | -105.1856
15.4984 | -105.1614
15.5021 | -105.1404
15.5058 | -105.122
15.5095 | -105.104
15.5131 | -105.0858
15.5168 | -105.0682
15.5205 | -105.0534
15.5242 | -105.0389
15.5279 | -105.0295
15.5316 | -105.0242
15.5353 | -105.0217
15.539 | -105.0207
15.5426 | -105.019
15.5463 | -105.0199
15.55 | -105.023
15.5537 | -105.0264
15.5574 | -105.0327
15.5611 | -105.0378
15.5648 | -105.0456
15.5685 | -105.0541
15.5722 | -105.0615
15.5758 | -105.0708
15.5795 | -105.0838
15.5832 | -105.0957
15.5869 | -105.1063
15.5906 | -105.1155
15.5943 | -105.1246
15.598 | -105.1316
15.6017 | -105.1361
15.6053 | -105.1395
15.609 | -105.1457
15.6127 | -105.15
15.6164 | -105.1535
15.6201 | -105.157
15.6238 | -105.1608
15.6275 | -105.1622
15.6312 | -105.1637
15.6349 | -105.1649
15.6385 | -105.1648
15.6422 | -105.1682
15.6459 | -105.1718
15.6496 | -105.176
15.6533 | -105.1781
15.657 | -105.1798
15.6607 | -105.1824
15.6644 | -105.1838
15.668 | -105.1915
15.6717 | -105.1992
15.6754 | -105.2041
15.6791 | -105.2114
15.6828 | -105.2185
15.6865 | -105.2251
15.6902 | -105.2332
15.6939 | -105.2418
15.6976 | -105.2512
15.7012 | -105.2557
15.7049 | -105.2573
15.7086 | -105.2578
15.7123 | -105.2578
15.716 | -105.2554
15.7197 | -105.2509
15.7234 | -105.2443
15.7271 | -105.2366
15.7308 | -105.2264
15.7344 | -105.2153
15.7381 | -105.2056
15.7418 | -105.1935
15.7455 | -105.1776
15.7492 | -105.1608
15.7529 | -105.1426
15.7566 | -105.1201
15.7603 | -105.0963
15.7639 | -105.0745
15.7676 | -105.0543
15.7713 | -105.0324
15.775 | -105.008
15.7787 | -104.9818
15.7824 | -104.9565
15.7861 | -104.9338
15.7898 | -104.9116
15.7935 | -104.8917
15.7971 | -104.8752
15.8008 | -104.858
15.8045 | -104.8413
15.8082 | -104.8269
15.8119 | -104.817
15.8156 | -104.8104
15.8193 | -104.8044
15.823 | -104.8019
15.8266 | -104.8015
15.8303 | -104.8015
15.834 | -104.8006
15.8377 | -104.806
15.8414 | -104.8148
15.8451 | -104.8238
15.8488 | -104.8357
15.8525 | -104.849
15.8562 | -104.8616
15.8598 | -104.8759
15.8635 | -104.8946
15.8672 | -104.9129
15.8709 | -104.9302
15.8746 | -104.9485
15.8783 | -104.9675
15.882 | -104.9864
15.8857 | -105.0091
15.8894 | -105.0316
15.893 | -105.049
15.8967 | -105.0671
15.9004 | -105.0871
15.9041 | -105.1046
15.9078 | -105.1207
15.9115 | -105.1349
15.9152 | -105.1462
15.9189 | -105.1603
15.9225 | -105.1718
15.9262 | -105.1838
15.9299 | -105.1993
15.9336 | -105.2147
15.9373 | -105.2282
15.941 | -105.2405
15.9447 | -105.2508
15.9484 | -105.2597
15.9521 | -105.2649
15.9557 | -105.2682
15.9594 | -105.2713
15.9631 | -105.2648
15.9668 | -105.2634
15.9705 | -105.2599
15.9742 | -105.2551
15.9779 | -105.2489
15.9816 | -105.2399
15.9852 | -105.2311
15.9889 | -105.2195
15.9926 | -105.2054
15.9963 | -105.1906
16 | -105.174
16.0037 | -105.1567
16.0074 | -105.1399
16.0111 | -105.1212
16.0148 | -105.0985
16.0184 | -105.0767
16.0221 | -105.0532
16.0258 | -105.0307
16.0295 | -105.0089
16.0332 | -104.9838
16.0369 | -104.9613
16.0406 | -104.9399
16.0443 | -104.9192
16.0479 | -104.9014
16.0516 | -104.8827
16.0553 | -104.8641
16.059 | -104.8487
16.0627 | -104.836
16.0664 | -104.8221
16.0701 | -104.8063
16.0738 | -104.7949
16.0775 | -104.7844
16.0811 | -104.7775
16.0848 | -104.775
16.0885 | -104.7737
16.0922 | -104.7702
16.0959 | -104.769
16.0996 | -104.7731
16.1033 | -104.7763
16.107 | -104.7798
16.1106 | -104.7865
16.1143 | -104.7949
16.118 | -104.8013
16.1217 | -104.8092
16.1254 | -104.8184
16.1291 | -104.8299
16.1328 | -104.8415
16.1365 | -104.8516
16.1402 | -104.8638
16.1438 | -104.8738
16.1475 | -104.8815
16.1512 | -104.8906
16.1549 | -104.8995
16.1586 | -104.9052
16.1623 | -104.9103
16.166 | -104.9154
16.1697 | -104.9209
16.1734 | -104.9255
16.177 | -104.9276
16.1807 | -104.931
16.1844 | -104.9344
16.1881 | -104.9365
16.1918 | -104.9352
16.1955 | -104.9334
16.1992 | -104.9319
16.2029 | -104.9298
16.2065 | -104.9276
16.2102 | -104.9262
16.2139 | -104.9247
16.2176 | -104.9211
16.2213 | -104.9186
16.225 | -104.9142
16.2287 | -104.9088
16.2324 | -104.903
16.2361 | -104.8998
16.2397 | -104.8967
16.2434 | -104.8928
16.2471 | -104.8916
16.2508 | -104.8898
16.2545 | -104.8894
16.2582 | -104.8877
16.2619 | -104.8869
16.2656 | -104.8903
16.2692 | -104.8922
16.2729 | -104.893
16.2766 | -104.8927
16.2803 | -104.8947
16.284 | -104.8994
16.2877 | -104.9045
16.2914 | -104.9097
16.2951 | -104.9161
16.2988 | -104.9244
16.3024 | -104.9358
16.3061 | -104.9477
16.3098 | -104.9613
16.3135 | -104.974
16.3172 | -104.9865
16.3209 | -104.9997
16.3246 | -105.0151
16.3283 | -105.0335
16.3319 | -105.0498
16.3356 | -105.0693
16.3393 | -105.0905
16.343 | -105.1083
16.3467 | -105.126
16.3504 | -105.1485
16.3541 | -105.1699
16.3578 | -105.1907
16.3615 | -105.2108
16.3651 | -105.2283
16.3688 | -105.2474
16.3725 | -105.2671
16.3762 | -105.2899
16.3799 | -105.3109
16.3836 | -105.3282
16.3873 | -105.3443
16.391 | -105.3608
16.3947 | -105.3735
16.3983 | -105.3887
16.402 | -105.4057
16.4057 | -105.4226
16.4094 | -105.4384
16.4131 | -105.4528
16.4168 | -105.4662
16.4205 | -105.4818
16.4242 | -105.4974
16.4278 | -105.5092
16.4315 | -105.523
16.4352 | -105.5395
16.4389 | -105.556
16.4426 | -105.5696
16.4463 | -105.5824
16.45 | -105.5961
16.4537 | -105.6099
16.4574 | -105.624
16.461 | -105.6395
16.4647 | -105.6566
16.4684 | -105.6676
16.4721 | -105.6789
16.4758 | -105.6927
16.4795 | -105.7078
16.4832 | -105.72
16.4869 | -105.7306
16.4905 | -105.7422
16.4942 | -105.7511
16.4979 | -105.7634
16.5016 | -105.7731
16.5053 | -105.7831
16.509 | -105.7939
16.5127 | -105.8071
16.5164 | -105.8155
16.5201 | -105.8221
16.5237 | -105.8307
16.5274 | -105.837
16.5311 | -105.8414
16.5348 | -105.8465
16.5385 | -105.8508
16.5422 | -105.8509
16.5459 | -105.8515
16.5496 | -105.8536
16.5532 | -105.8563
16.5569 | -105.8575
16.5606 | -105.8568
16.5643 | -105.8517
16.568 | -105.8457
16.5717 | -105.8395
16.5754 | -105.8327
16.5791 | -105.8274
16.5828 | -105.8213
16.5864 | -105.813
16.5901 | -105.8062
16.5938 | -105.8004
16.5975 | -105.7952
16.6012 | -105.7919
16.6049 | -105.7883
16.6086 | -105.7848
16.6123 | -105.7785
16.616 | -105.773
16.6196 | -105.77
16.6233 | -105.7655
16.627 | -105.7608
16.6307 | -105.7588
16.6344 | -105.7538
16.6381 | -105.7501
16.6418 | -105.7472
16.6455 | -105.7463
16.6491 | -105.7456
16.6528 | -105.7459
16.6565 | -105.7406
16.6602 | -105.7387
16.6639 | -105.7383
16.6676 | -105.7339
16.6713 | -105.7274
16.675 | -105.722
16.6787 | -105.7183
16.6823 | -105.714
16.686 | -105.7099
16.6897 | -105.7037
16.6934 | -105.6985
16.6971 | -105.6915
16.7008 | -105.6839
16.7045 | -105.6808
16.7082 | -105.6773
16.7118 | -105.6719
16.7155 | -105.6633
16.7192 | -105.6578
16.7229 | -105.6529
16.7266 | -105.6433
16.7303 | -105.6343
16.734 | -105.6236
16.7377 | -105.6137
16.7414 | -105.6061
16.745 | -105.597
16.7487 | -105.5877
16.7524 | -105.5792
16.7561 | -105.5735
16.7598 | -105.5675
16.7635 | -105.5607
16.7672 | -105.5553
16.7709 | -105.55
16.7746 | -105.5419
16.7782 | -105.5353
16.7819 | -105.531
16.7856 | -105.5252
16.7893 | -105.5189
16.793 | -105.5136
16.7967 | -105.5101
16.8004 | -105.5058
16.8041 | -105.5055
16.8077 | -105.5067
16.8114 | -105.5069
16.8151 | -105.506
16.8188 | -105.5066
16.8225 | -105.5078
16.8262 | -105.5092
16.8299 | -105.5114
16.8336 | -105.5101
16.8373 | -105.5088
16.8409 | -105.5081
16.8446 | -105.5088
16.8483 | -105.5089
16.852 | -105.5099
16.8557 | -105.5113
16.8594 | -105.5136
16.8631 | -105.5191
16.8668 | -105.5244
16.8704 | -105.5274
16.8741 | -105.5308
16.8778 | -105.5336
16.8815 | -105.5361
16.8852 | -105.5397
16.8889 | -105.5425
16.8926 | -105.5437
16.8963 | -105.5476
16.9 | -105.5494
16.9036 | -105.5511
16.9073 | -105.5523
16.911 | -105.5515
16.9147 | -105.5509
16.9184 | -105.5501
16.9221 | -105.5495
16.9258 | -105.5506
16.9295 | -105.5498
16.9331 | -105.5461
16.9368 | -105.5424
16.9405 | -105.5389
16.9442 | -105.5318
16.9479 | -105.5246
16.9516 | -105.5168
16.9553 | -105.5112
16.959 | -105.5032
16.9627 | -105.4964
16.9663 | -105.4914
16.97 | -105.4808
16.9737 | -105.4688
16.9774 | -105.4553
16.9811 | -105.4397
16.9848 | -105.4245
16.9885 | -105.4101
16.9922 | -105.3953
16.9959 | -105.38
16.9995 | -105.3669
17.0032 | -105.3522
17.0069 | -105.3339
17.0106 | -105.3188
17.0143 | -105.308
17.018 | -105.2978
17.0217 | -105.2878
17.0254 | -105.2867
17.029 | -105.2801
17.0327 | -105.2745
17.0364 | -105.269
17.0401 | -105.2642
17.0438 | -105.2627
17.0475 | -105.2612
17.0512 | -105.2565
17.0549 | -105.2531
17.0586 | -105.2541
17.0622 | -105.2526
17.0659 | -105.2528
17.0696 | -105.2533
17.0733 | -105.255
17.077 | -105.2587
17.0807 | -105.2593
17.0844 | -105.2623
17.0881 | -105.2655
17.0917 | -105.2664
17.0954 | -105.2668
17.0991 | -105.2667
17.1028 | -105.264
17.1065 | -105.2602
17.1102 | -105.2565
17.1139 | -105.2529
17.1176 | -105.249
17.1213 | -105.2436
17.1249 | -105.2374
17.1286 | -105.2295
17.1323 | -105.2225
17.136 | -105.2178
17.1397 | -105.2121
17.1434 | -105.2061
17.1471 | -105.198
17.1508 | -105.1905
17.1544 | -105.1852
17.1581 | -105.1789
17.1618 | -105.1703
17.1655 | -105.1645
17.1692 | -105.1582
17.1729 | -105.1511
17.1766 | -105.1448
17.1803 | -105.1397
17.184 | -105.1337
17.1876 | -105.1322
17.1913 | -105.1316
17.195 | -105.1302
17.1987 | -105.1275
17.2024 | -105.1273
17.2061 | -105.1251
17.2098 | -105.1238
17.2135 | -105.1252
17.2172 | -105.1248
17.2208 | -105.1261
17.2245 | -105.1287
17.2282 | -105.1316
17.2319 | -105.1334
17.2356 | -105.1344
17.2393 | -105.1335
17.243 | -105.1311
17.2467 | -105.1314
17.2503 | -105.1342
17.254 | -105.1336
17.2577 | -105.1306
17.2614 | -105.1278
17.2651 | -105.1238
17.2688 | -105.1195
17.2725 | -105.1163
17.2762 | -105.1141
17.2799 | -105.1119
17.2835 | -105.11
17.2872 | -105.1064
17.2909 | -105.1015
17.2946 | -105.0972
17.2983 | -105.0971
17.302 | -105.0967
17.3057 | -105.0925
17.3094 | -105.0894
17.313 | -105.0869
17.3167 | -105.0872
17.3204 | -105.0853
17.3241 | -105.0855
17.3278 | -105.0885
17.3315 | -105.0891
17.3352 | -105.0898
17.3389 | -105.0929
17.3426 | -105.0931
17.3462 | -105.0947
17.3499 | -105.0974
17.3536 | -105.0985
17.3573 | -105.1008
17.361 | -105.1032
17.3647 | -105.1022
17.3684 | -105.0993
17.3721 | -105.0967
17.3757 | -105.0942
17.3794 | -105.0913
17.3831 | -105.0852
17.3868 | -105.0793
17.3905 | -105.0739
17.3942 | -105.0672
17.3979 | -105.0605
17.4016 | -105.0518
17.4053 | -105.0416
17.4089 | -105.0335
17.4126 | -105.0196
17.4163 | -105.0023
17.42 | -104.9874
17.4237 | -104.9733
17.4274 | -104.9559
17.4311 | -104.9384
17.4348 | -104.9167
17.4385 | -104.8942
17.4421 | -104.8717
17.4458 | -104.8507
17.4495 | -104.8272
17.4532 | -104.8064
17.4569 | -104.7862
17.4606 | -104.7655
17.4643 | -104.7461
17.468 | -104.7261
17.4716 | -104.705
17.4753 | -104.6825
17.479 | -104.6597
17.4827 | -104.6377
17.4864 | -104.6207
17.4901 | -104.6065
17.4938 | -104.592
17.4975 | -104.5777
17.5012 | -104.5644
17.5048 | -104.5513
17.5085 | -104.5379
17.5122 | -104.5259
17.5159 | -104.5143
17.5196 | -104.5051
17.5233 | -104.4964
17.527 | -104.4881
17.5307 | -104.4806
17.5343 | -104.4742
17.538 | -104.4683
17.5417 | -104.4595
17.5454 | -104.4514
17.5491 | -104.4454
17.5528 | -104.4429
17.5565 | -104.4403
17.5602 | -104.4381
17.5639 | -104.4337
17.5675 | -104.4294
17.5712 | -104.427
17.5749 | -104.4249
17.5786 | -104.4212
17.5823 | -104.4187
17.586 | -104.4162
17.5897 | -104.4148
17.5934 | -104.4156
17.597 | -104.414
17.6007 | -104.4126
17.6044 | -104.4126
17.6081 | -104.4136
17.6118 | -104.4162
17.6155 | -104.4211
17.6192 | -104.4265
17.6229 | -104.4321
17.6266 | -104.4368
17.6302 | -104.4377
17.6339 | -104.439
17.6376 | -104.4426
17.6413 | -104.4454
17.645 | -104.4468
17.6487 | -104.4461
17.6524 | -104.4465
17.6561 | -104.4469
17.6598 | -104.4455
17.6634 | -104.4439
17.6671 | -104.4449
17.6708 | -104.4433
17.6745 | -104.4397
17.6782 | -104.4375
17.6819 | -104.4321
17.6856 | -104.4252
17.6893 | -104.4189
17.6929 | -104.4143
17.6966 | -104.4079
17.7003 | -104.4007
17.704 | -104.3936
17.7077 | -104.3872
17.7114 | -104.3805
17.7151 | -104.3723
17.7188 | -104.3642
17.7225 | -104.3576
17.7261 | -104.351
17.7298 | -104.3461
17.7335 | -104.3436
17.7372 | -104.3396
17.7409 | -104.3363
17.7446 | -104.3339
17.7483 | -104.33
17.752 | -104.3269
17.7556 | -104.3259
17.7593 | -104.3262
17.763 | -104.3272
17.7667 | -104.3285
17.7704 | -104.3275
17.7741 | -104.3267
17.7778 | -104.3279
17.7815 | -104.3303
17.7852 | -104.3324
17.7888 | -104.3349
17.7925 | -104.3391
17.7962 | -104.3434
17.7999 | -104.3466
17.8036 | -104.349
17.8073 | -104.3511
17.811 | -104.3521
17.8147 | -104.3542
17.8183 | -104.357
17.822 | -104.3594
17.8257 | -104.3616
17.8294 | -104.3613
17.8331 | -104.3615
17.8368 | -104.3625
17.8405 | -104.3638
17.8442 | -104.362
17.8479 | -104.3618
17.8515 | -104.3623
17.8552 | -104.3614
17.8589 | -104.3578
17.8626 | -104.3515
17.8663 | -104.3446
17.87 | -104.3384
17.8737 | -104.3331
17.8774 | -104.3261
17.8811 | -104.3189
17.8847 | -104.3139
17.8884 | -104.3053
17.8921 | -104.292
17.8958 | -104.2799
17.8995 | -104.2696
17.9032 | -104.2591
17.9069 | -104.2438
17.9106 | -104.2273
17.9142 | -104.2099
17.9179 | -104.1929
17.9216 | -104.174
17.9253 | -104.1533
17.929 | -104.1315
17.9327 | -104.1068
17.9364 | -104.0823
17.9401 | -104.0583
17.9438 | -104.0335
17.9474 | -104.0086
17.9511 | -103.9841
17.9548 | -103.9598
17.9585 | -103.9328
17.9622 | -103.9034
17.9659 | -103.8768
17.9696 | -103.853
17.9733 | -103.8258
17.9769 | -103.7992
17.9806 | -103.7736
17.9843 | -103.7477
17.988 | -103.7187
17.9917 | -103.6902
17.9954 | -103.6589
17.9991 | -103.6307
18.0028 | -103.6065
18.0065 | -103.5822
18.0101 | -103.5578
18.0138 | -103.5344
18.0175 | -103.5086
18.0212 | -103.4821
18.0249 | -103.4588
18.0286 | -103.436
18.0323 | -103.4147
18.036 | -103.3919
18.0397 | -103.3716
18.0433 | -103.3501
18.047 | -103.3285
18.0507 | -103.3062
18.0544 | -103.285
18.0581 | -103.2655
18.0618 | -103.245
18.0655 | -103.2256
18.0692 | -103.2063
18.0728 | -103.1878
18.0765 | -103.1678
18.0802 | -103.1477
18.0839 | -103.1304
18.0876 | -103.1158
18.0913 | -103.1014
18.095 | -103.0854
18.0987 | -103.0651
18.1024 | -103.0471
18.106 | -103.0299
18.1097 | -103.0145
18.1134 | -102.9981
18.1171 | -102.9823
18.1208 | -102.9683
18.1245 | -102.9536
18.1282 | -102.9384
18.1319 | -102.9247
18.1355 | -102.9115
18.1392 | -102.8995
18.1429 | -102.8891
18.1466 | -102.8797
18.1503 | -102.8695
18.154 | -102.859
18.1577 | -102.8493
18.1614 | -102.8415
18.1651 | -102.8373
18.1687 | -102.8311
18.1724 | -102.8263
18.1761 | -102.8243
18.1798 | -102.8223
18.1835 | -102.8208
18.1872 | -102.8212
18.1909 | -102.8239
18.1946 | -102.8267
18.1982 | -102.8283
18.2019 | -102.8296
18.2056 | -102.8322
18.2093 | -102.8358
18.213 | -102.84
18.2167 | -102.845
18.2204 | -102.8527
18.2241 | -102.8587
18.2278 | -102.8646
18.2314 | -102.8717
18.2351 | -102.8766
18.2388 | -102.8847
18.2425 | -102.8923
18.2462 | -102.9012
18.2499 | -102.9093
18.2536 | -102.9207
18.2573 | -102.9293
18.261 | -102.9382
18.2646 | -102.9475
18.2683 | -102.9563
18.272 | -102.9653
18.2757 | -102.9773
18.2794 | -102.9888
18.2831 | -102.9984
18.2868 | -103.0089
18.2905 | -103.0183
18.2941 | -103.0279
18.2978 | -103.0357
18.3015 | -103.0456
18.3052 | -103.0568
18.3089 | -103.0688
18.3126 | -103.0791
18.3163 | -103.0869
18.32 | -103.0976
18.3237 | -103.1073
18.3273 | -103.1161
18.331 | -103.1248
18.3347 | -103.134
18.3384 | -103.1419
18.3421 | -103.1478
18.3458 | -103.1512
18.3495 | -103.1557
18.3532 | -103.1579
18.3568 | -103.1615
18.3605 | -103.1704
18.3642 | -103.1777
18.3679 | -103.18
18.3716 | -103.1799
18.3753 | -103.1807
18.379 | -103.1843
18.3827 | -103.1869
18.3864 | -103.1902
18.39 | -103.1915
18.3937 | -103.1878
18.3974 | -103.1856
18.4011 | -103.1849
18.4048 | -103.1841
18.4085 | -103.1841
18.4122 | -103.1865
18.4159 | -103.1874
18.4195 | -103.1872
18.4232 | -103.1839
18.4269 | -103.1797
18.4306 | -103.1783
18.4343 | -103.1769
18.438 | -103.1718
18.4417 | -103.1637
18.4454 | -103.1545
18.4491 | -103.1468
18.4527 | -103.1361
18.4564 | -103.1275
18.4601 | -103.1199
18.4638 | -103.1106
18.4675 | -103.0988
18.4712 | -103.0878
18.4749 | -103.0752
18.4786 | -103.0605
18.4823 | -103.0455
18.4859 | -103.0309
18.4896 | -103.016
18.4933 | -103.0023
18.497 | -102.9913
18.5007 | -102.978
18.5044 | -102.9642
18.5081 | -102.9519
18.5118 | -102.9398
18.5154 | -102.9267
18.5191 | -102.9148
18.5228 | -102.9047
18.5265 | -102.8951
18.5302 | -102.8867
18.5339 | -102.8793
18.5376 | -102.8727
18.5413 | -102.8677
18.545 | -102.8619
18.5486 | -102.8569
18.5523 | -102.8549
18.556 | -102.8566
18.5597 | -102.8591
18.5634 | -102.8614
18.5671 | -102.8635
18.5708 | -102.8674
18.5745 | -102.8726
18.5781 | -102.8801
18.5818 | -102.8874
18.5855 | -102.8932
18.5892 | -102.8982
18.5929 | -102.9046
18.5966 | -102.9133
18.6003 | -102.9233
18.604 | -102.9324
18.6077 | -102.9418
18.6113 | -102.9517
18.615 | -102.9642
18.6187 | -102.9736
18.6224 | -102.9779
18.6261 | -102.987
18.6298 | -102.9964
18.6335 | -103.007
18.6372 | -103.0166
18.6408 | -103.0284
18.6445 | -103.037
18.6482 | -103.0452
18.6519 | -103.0518
18.6556 | -103.0613
18.6593 | -103.0721
18.663 | -103.0793
18.6667 | -103.084
18.6704 | -103.0882
18.674 | -103.0923
18.6777 | -103.0966
18.6814 | -103.101
18.6851 | -103.1046
18.6888 | -103.1091
18.6925 | -103.1133
18.6962 | -103.1147
18.6999 | -103.1161
18.7035 | -103.117
18.7072 | -103.1203
18.7109 | -103.1244
18.7146 | -103.1244
18.7183 | -103.1255
18.722 | -103.1253
18.7257 | -103.1244
18.7294 | -103.1235
18.7331 | -103.1225
18.7367 | -103.1199
18.7404 | -103.1191
18.7441 | -103.1205
18.7478 | -103.1203
18.7515 | -103.1162
18.7552 | -103.1125
18.7589 | -103.1092
18.7626 | -103.102
18.7663 | -103.0964
18.7699 | -103.0942
18.7736 | -103.0883
18.7773 | -103.0824
18.781 | -103.0765
18.7847 | -103.0708
18.7884 | -103.0641
18.7921 | -103.0595
18.7958 | -103.054
18.7994 | -103.0467
18.8031 | -103.0393
18.8068 | -103.0303
18.8105 | -103.0219
18.8142 | -103.0135
18.8179 | -103.0026
18.8216 | -102.991
18.8253 | -102.9815
18.829 | -102.9708
18.8326 | -102.9598
18.8363 | -102.9497
18.84 | -102.9393
18.8437 | -102.9289
18.8474 | -102.9159
18.8511 | -102.9038
18.8548 | -102.8934
18.8585 | -102.8813
18.8621 | -102.8669
18.8658 | -102.8554
18.8695 | -102.844
18.8732 | -102.832
18.8769 | -102.8181
18.8806 | -102.8053
18.8843 | -102.7945
18.888 | -102.7863
18.8917 | -102.7771
18.8953 | -102.7662
18.899 | -102.7553
18.9027 | -102.7468
18.9064 | -102.7403
18.9101 | -102.7308
18.9138 | -102.7243
18.9175 | -102.7186
18.9212 | -102.7148
18.9248 | -102.7106
18.9285 | -102.7078
18.9322 | -102.7065
18.9359 | -102.7062
18.9396 | -102.7056
18.9433 | -102.7051
18.947 | -102.7042
18.9507 | -102.704
18.9544 | -102.7073
18.958 | -102.7099
18.9617 | -102.7135
18.9654 | -102.7181
18.9691 | -102.7216
18.9728 | -102.7284
18.9765 | -102.7342
18.9802 | -102.7381
18.9839 | -102.7426
18.9876 | -102.7467
18.9912 | -102.7518
18.9949 | -102.7558
18.9986 | -102.7608
19.0023 | -102.7676
19.006 | -102.7743
19.0097 | -102.7787
19.0134 | -102.7823
19.0171 | -102.7876
19.0207 | -102.7975
19.0244 | -102.8066
19.0281 | -102.8143
19.0318 | -102.8238
19.0355 | -102.8284
19.0392 | -102.8314
19.0429 | -102.8395
19.0466 | -102.848
19.0503 | -102.8541
19.0539 | -102.8608
19.0576 | -102.8672
19.0613 | -102.8712
19.065 | -102.876
19.0687 | -102.8851
19.0724 | -102.8915
19.0761 | -102.8975
19.0798 | -102.9045
19.0834 | -102.9078
19.0871 | -102.9094
19.0908 | -102.9128
19.0945 | -102.9158
19.0982 | -102.9184
19.1019 | -102.9198
19.1056 | -102.9189
19.1093 | -102.9158
19.113 | -102.9109
19.1166 | -102.9078
19.1203 | -102.9037
19.124 | -102.8957
19.1277 | -102.8887
19.1314 | -102.8819
19.1351 | -102.8722
19.1388 | -102.8617
19.1425 | -102.8486
19.1461 | -102.8364
19.1498 | -102.8235
19.1535 | -102.8064
19.1572 | -102.7921
19.1609 | -102.7777
19.1646 | -102.7612
19.1683 | -102.7448
19.172 | -102.7311
19.1757 | -102.7174
19.1793 | -102.7043
19.183 | -102.6926
19.1867 | -102.6798
19.1904 | -102.6686
19.1941 | -102.6566
19.1978 | -102.6447
19.2015 | -102.6349
19.2052 | -102.6274
19.2089 | -102.6207
19.2125 | -102.6143
19.2162 | -102.6093
19.2199 | -102.6044
19.2236 | -102.602
19.2273 | -102.5995
19.231 | -102.6001
19.2347 | -102.6004
19.2384 | -102.6042
19.242 | -102.6094
19.2457 | -102.6129
19.2494 | -102.6193
19.2531 | -102.6308
19.2568 | -102.645
19.2605 | -102.6606
19.2642 | -102.6751
19.2679 | -102.6926
19.2716 | -102.7108
19.2752 | -102.7276
19.2789 | -102.747
19.2826 | -102.7685
19.2863 | -102.7907
19.29 | -102.8122
19.2937 | -102.8324
19.2974 | -102.8536
19.3011 | -102.8758
19.3047 | -102.8988
19.3084 | -102.9236
19.3121 | -102.9477
19.3158 | -102.9693
19.3195 | -102.9908
19.3232 | -103.0139
19.3269 | -103.0355
19.3306 | -103.056
19.3343 | -103.0767
19.3379 | -103.0974
19.3416 | -103.114
19.3453 | -103.1316
19.349 | -103.1492
19.3527 | -103.1652
19.3564 | -103.1808
19.3601 | -103.197
19.3638 | -103.2081
19.3675 | -103.2183
19.3711 | -103.2294
19.3748 | -103.2406
19.3785 | -103.2529
19.3822 | -103.2666
19.3859 | -103.2804
19.3896 | -103.2949
19.3933 | -103.3093
19.397 | -103.3235
19.4006 | -103.3365
19.4043 | -103.3499
19.408 | -103.3601
19.4117 | -103.3728
19.4154 | -103.3893
19.4191 | -103.4035
19.4228 | -103.4184
19.4265 | -103.4313
19.4302 | -103.4452
19.4338 | -103.4599
19.4375 | -103.4729
19.4412 | -103.4876
19.4449 | -103.5024
19.4486 | -103.5163
19.4523 | -103.5295
19.456 | -103.5421
19.4597 | -103.5547
19.4633 | -103.5674
19.467 | -103.5804
19.4707 | -103.5951
19.4744 | -103.6107
19.4781 | -103.6236
19.4818 | -103.6341
19.4855 | -103.6436
19.4892 | -103.6542
19.4929 | -103.6662
19.4965 | -103.6788
19.5002 | -103.6902
19.5039 | -103.7003
19.5076 | -103.7089
19.5113 | -103.7176
19.515 | -103.7266
19.5187 | -103.7343
19.5224 | -103.7422
19.526 | -103.7498
19.5297 | -103.7561
19.5334 | -103.7582
19.5371 | -103.7619
19.5408 | -103.7661
19.5445 | -103.7687
19.5482 | -103.771
19.5519 | -103.7733
19.5556 | -103.773
19.5592 | -103.7734
19.5629 | -103.7728
19.5666 | -103.7677
19.5703 | -103.7618
19.574 | -103.7565
19.5777 | -103.7513
19.5814 | -103.7447
19.5851 | -103.7383
19.5888 | -103.731
19.5924 | -103.7225
19.5961 | -103.7142
19.5998 | -103.7037
19.6035 | -103.6894
19.6072 | -103.677
19.6109 | -103.6661
19.6146 | -103.6512
19.6183 | -103.6366
19.6219 | -103.6246
19.6256 | -103.6126
19.6293 | -103.5995
19.633 | -103.5865
19.6367 | -103.5732
19.6404 | -103.5614
19.6441 | -103.5469
19.6478 | -103.5302
19.6515 | -103.5142
19.6551 | -103.4972
19.6588 | -103.4808
19.6625 | -103.4644
19.6662 | -103.4479
19.6699 | -103.4307
19.6736 | -103.4145
19.6773 | -103.3972
19.681 | -103.378
19.6846 | -103.3624
19.6883 | -103.348
19.692 | -103.3322
19.6957 | -103.316
19.6994 | -103.3004
19.7031 | -103.2852
19.7068 | -103.2721
19.7105 | -103.2618
19.7142 | -103.2514
19.7178 | -103.2406
19.7215 | -103.2301
19.7252 | -103.2196
19.7289 | -103.2113
19.7326 | -103.2046
19.7363 | -103.1998
19.74 | -103.196
19.7437 | -103.1932
19.7473 | -103.1931
19.751 | -103.1941
19.7547 | -103.1964
19.7584 | -103.1998
19.7621 | -103.2042
19.7658 | -103.2087
19.7695 | -103.2135
19.7732 | -103.2188
19.7769 | -103.2262
19.7805 | -103.2344
19.7842 | -103.244
19.7879 | -103.253
19.7916 | -103.2631
19.7953 | -103.2721
19.799 | -103.2816
19.8027 | -103.2935
19.8064 | -103.3047
19.8101 | -103.3149
19.8137 | -103.324
19.8174 | -103.332
19.8211 | -103.341
19.8248 | -103.3502
19.8285 | -103.3567
19.8322 | -103.3634
19.8359 | -103.3716
19.8396 | -103.3793
19.8432 | -103.3864
19.8469 | -103.3947
19.8506 | -103.4011
19.8543 | -103.4055
19.858 | -103.4078
19.8617 | -103.4116
19.8654 | -103.4121
19.8691 | -103.4107
19.8728 | -103.4094
19.8764 | -103.4072
19.8801 | -103.4062
19.8838 | -103.4074
19.8875 | -103.406
19.8912 | -103.403
19.8949 | -103.4005
19.8986 | -103.3972
19.9023 | -103.3924
19.9059 | -103.3872
19.9096 | -103.3811
19.9133 | -103.3757
19.917 | -103.3717
19.9207 | -103.366
19.9244 | -103.362
19.9281 | -103.3591
19.9318 | -103.3527
19.9355 | -103.3468
19.9391 | -103.3416
19.9428 | -103.3369
19.9465 | -103.33
19.9502 | -103.3245
19.9539 | -103.3187
19.9576 | -103.3152
19.9613 | -103.312
19.965 | -103.3058
19.9686 | -103.2973
19.9723 | -103.2916
19.976 | -103.2863
19.9797 | -103.2791
19.9834 | -103.2744
19.9871 | -103.2679
19.9908 | -103.2577
19.9945 | -103.2488
19.9982 | -103.2424
20.0018 | -103.2375
20.0055 | -103.2324
20.0092 | -103.2289
20.0129 | -103.2231
20.0166 | -103.2132
20.0203 | -103.2049
20.024 | -103.1981
20.0277 | -103.1914
20.0314 | -103.1817
20.035 | -103.1714
20.0387 | -103.1619
20.0424 | -103.1525
20.0461 | -103.1438
20.0498 | -103.1359
20.0535 | -103.1292
20.0572 | -103.1219
20.0609 | -103.1129
20.0645 | -103.1045
20.0682 | -103.0942
20.0719 | -103.085
20.0756 | -103.0771
20.0793 | -103.0668
20.083 | -103.0561
20.0867 | -103.0435
20.0904 | -103.0318
20.0941 | -103.0212
20.0977 | -103.0127
20.1014 | -103.0046
20.1051 | -102.9961
20.1088 | -102.9876
20.1125 | -102.9814
20.1162 | -102.975
20.1199 | -102.967
20.1236 | -102.9608
20.1272 | -102.9557
20.1309 | -102.9523
20.1346 | -102.9522
20.1383 | -102.9516
20.142 | -102.953
20.1457 | -102.9558
20.1494 | -102.9618
20.1531 | -102.97
20.1568 | -102.9781
20.1604 | -102.9864
20.1641 | -102.9952
20.1678 | -103.003
20.1715 | -103.0115
20.1752 | -103.0223
20.1789 | -103.0329
20.1826 | -103.0425
20.1863 | -103.0528
20.1899 | -103.0635
20.1936 | -103.074
20.1973 | -103.0852
20.201 | -103.0951
20.2047 | -103.105
20.2084 | -103.1157
20.2121 | -103.127
20.2158 | -103.1355
20.2195 | -103.1385
20.2231 | -103.1416
20.2268 | -103.1446
20.2305 | -103.1483
20.2342 | -103.1523
20.2379 | -103.1547
20.2416 | -103.1568
20.2453 | -103.1566
20.249 | -103.1553
20.2527 | -103.1569
20.2563 | -103.1576
20.26 | -103.154
20.2637 | -103.1494
20.2674 | -103.1452
20.2711 | -103.1406
20.2748 | -103.1353
20.2785 | -103.13
20.2822 | -103.1199
20.2858 | -103.1084
20.2895 | -103.0972
20.2932 | -103.0866
20.2969 | -103.0751
20.3006 | -103.0635
20.3043 | -103.0534
20.308 | -103.0411
20.3117 | -103.0288
20.3154 | -103.0167
20.319 | -103.0039
20.3227 | -102.9924
20.3264 | -102.9771
20.3301 | -102.9635
20.3338 | -102.9469
20.3375 | -102.9308
20.3412 | -102.9165
20.3449 | -102.9026
20.3485 | -102.888
20.3522 | -102.8748
20.3559 | -102.8622
20.3596 | -102.8499
20.3633 | -102.8377
20.367 | -102.827
20.3707 | -102.8183
20.3744 | -102.8102
20.3781 | -102.8003
20.3817 | -102.7916
20.3854 | -102.7849
20.3891 | -102.7781
20.3928 | -102.7671
20.3965 | -102.7573
20.4002 | -102.7512
20.4039 | -102.7443
20.4076 | -102.737
20.4112 | -102.7257
20.4149 | -102.7129
20.4186 | -102.7049
20.4223 | -102.6979
20.426 | -102.6902
20.4297 | -102.6811
20.4334 | -102.6725
20.4371 | -102.6648
20.4408 | -102.6539
20.4444 | -102.6434
20.4481 | -102.635
20.4518 | -102.6234
20.4555 | -102.6125
20.4592 | -102.5986
20.4629 | -102.5848
20.4666 | -102.5722
20.4703 | -102.5599
20.474 | -102.5444
20.4776 | -102.5262
20.4813 | -102.5047
20.485 | -102.4809
20.4887 | -102.4566
20.4924 | -102.4321
20.4961 | -102.4055
20.4998 | -102.3794
20.5035 | -102.3563
20.5071 | -102.3291
20.5108 | -102.2981
20.5145 | -102.2693
20.5182 | -102.2393
20.5219 | -102.2047
20.5256 | -102.1693
20.5293 | -102.1355
20.533 | -102.0993
20.5367 | -102.0603
20.5403 | -102.0211
20.544 | -101.9815
20.5477 | -101.9389
20.5514 | -101.8959
20.5551 | -101.8554
20.5588 | -101.8154
20.5625 | -101.7698
20.5662 | -101.7252
20.5698 | -101.6858
20.5735 | -101.6461
20.5772 | -101.6072
20.5809 | -101.5683
20.5846 | -101.527
20.5883 | -101.4844
20.592 | -101.4416
20.5957 | -101.3951
20.5994 | -101.3494
20.603 | -101.3038
20.6067 | -101.2536
20.6104 | -101.2068
20.6141 | -101.1618
20.6178 | -101.1196
20.6215 | -101.0749
20.6252 | -101.0253
20.6289 | -100.9741
20.6325 | -100.9197
20.6362 | -100.8689
20.6399 | -100.8223
20.6436 | -100.7745
20.6473 | -100.7255
20.651 | -100.6769
20.6547 | -100.6291
20.6584 | -100.5835
20.6621 | -100.5347
20.6657 | -100.4857
20.6694 | -100.441
20.6731 | -100.3956
20.6768 | -100.3473
20.6805 | -100.3007
20.6842 | -100.2588
20.6879 | -100.2178
20.6916 | -100.1744
20.6953 | -100.1337
20.6989 | -100.0963
20.7026 | -100.061
20.7063 | -100.0259
20.71 | -99.9878
20.7137 | -99.9471
20.7174 | -99.9102
20.7211 | -99.8764
20.7248 | -99.8439
20.7284 | -99.8145
20.7321 | -99.7879
20.7358 | -99.7623
20.7395 | -99.7379
20.7432 | -99.7187
20.7469 | -99.701
20.7506 | -99.697
20.7543 | -99.6809
20.758 | -99.6642
20.7616 | -99.6507
20.7653 | -99.6397
20.769 | -99.6307
20.7727 | -99.6219
20.7764 | -99.6163
20.7801 | -99.6158
20.7838 | -99.6143
20.7875 | -99.6114
20.7911 | -99.6088
20.7948 | -99.61
20.7985 | -99.6138
20.8022 | -99.6176
20.8059 | -99.6244
20.8096 | -99.6312
20.8133 | -99.6374
20.817 | -99.6456
20.8207 | -99.6535
20.8243 | -99.6594
20.828 | -99.6673
20.8317 | -99.6752
20.8354 | -99.6841
20.8391 | -99.6937
20.8428 | -99.7028
20.8465 | -99.7115
20.8502 | -99.7186
20.8539 | -99.7256
20.8575 | -99.7338
20.8612 | -99.7395
20.8649 | -99.7428
20.8686 | -99.7465
20.8723 | -99.7505
20.876 | -99.7541
20.8797 | -99.7579
20.8834 | -99.7608
20.887 | -99.7639
20.8907 | -99.7666
20.8944 | -99.7675
20.8981 | -99.7686
20.9018 | -99.7734
20.9055 | -99.7752
20.9092 | -99.7769
20.9129 | -99.7766
20.9166 | -99.7788
20.9202 | -99.7846
20.9239 | -99.7901
20.9276 | -99.7955
20.9313 | -99.7997
20.935 | -99.8015
20.9387 | -99.8035
20.9424 | -99.8096
20.9461 | -99.8162
20.9497 | -99.8217
20.9534 | -99.8263
20.9571 | -99.8331
20.9608 | -99.8387
20.9645 | -99.8429
20.9682 | -99.8496
20.9719 | -99.8592
20.9756 | -99.8666
20.9793 | -99.8725
20.9829 | -99.8784
20.9866 | -99.8865
20.9903 | -99.8933
20.994 | -99.8992
20.9977 | -99.9066
21.0014 | -99.9151
21.0051 | -99.9244
21.0088 | -99.9367
21.0124 | -99.9459
21.0161 | -99.9528
21.0198 | -99.9595
21.0235 | -99.9669
21.0272 | -99.9731
21.0309 | -99.9788
21.0346 | -99.9849
21.0383 | -99.9896
21.042 | -99.9938
21.0456 | -99.9986
21.0493 | -100.0072
21.053 | -100.0169
21.0567 | -100.0265
21.0604 | -100.0352
21.0641 | -100.0433
21.0678 | -100.0521
21.0715 | -100.0654
21.0752 | -100.0778
21.0788 | -100.0857
21.0825 | -100.0966
21.0862 | -100.1108
21.0899 | -100.1265
21.0936 | -100.1386
21.0973 | -100.1489
21.101 | -100.1618
21.1047 | -100.1751
21.1083 | -100.192
21.112 | -100.2101
21.1157 | -100.2287
21.1194 | -100.2478
21.1231 | -100.2679
21.1268 | -100.2891
21.1305 | -100.3117
21.1342 | -100.3334
21.1379 | -100.3531
21.1415 | -100.3724
21.1452 | -100.3924
21.1489 | -100.4119
21.1526 | -100.4319
21.1563 | -100.4536
21.16 | -100.4766
21.1637 | -100.4979
21.1674 | -100.5178
21.171 | -100.535
21.1747 | -100.552
21.1784 | -100.5687
21.1821 | -100.5862
21.1858 | -100.6048
21.1895 | -100.624
21.1932 | -100.6416
21.1969 | -100.6553
21.2006 | -100.6664
21.2042 | -100.6777
21.2079 | -100.6869
21.2116 | -100.6956
21.2153 | -100.7014
21.219 | -100.7077
21.2227 | -100.7126
21.2264 | -100.7164
21.2301 | -100.7175
21.2337 | -100.7185
21.2374 | -100.7207
21.2411 | -100.7189
21.2448 | -100.717
21.2485 | -100.7153
21.2522 | -100.7133
21.2559 | -100.7129
21.2596 | -100.7147
21.2633 | -100.714
21.2669 | -100.7102
21.2706 | -100.7067
21.2743 | -100.7062
21.278 | -100.7058
21.2817 | -100.7061
21.2854 | -100.7059
21.2891 | -100.7036
21.2928 | -100.7032
21.2965 | -100.7039
21.3001 | -100.704
21.3038 | -100.7055
21.3075 | -100.7076
21.3112 | -100.7104
21.3149 | -100.715
21.3186 | -100.7208
21.3223 | -100.7254
21.326 | -100.7266
21.3296 | -100.7333
21.3333 | -100.7402
21.337 | -100.7488
21.3407 | -100.758
21.3444 | -100.7678
21.3481 | -100.7782
21.3518 | -100.7865
21.3555 | -100.7933
21.3592 | -100.8008
21.3628 | -100.8063
21.3665 | -100.8125
21.3702 | -100.8189
21.3739 | -100.823
21.3776 | -100.8301
21.3813 | -100.8365
21.385 | -100.8415
21.3887 | -100.8478
21.3923 | -100.8519
21.396 | -100.8544
21.3997 | -100.8585
21.4034 | -100.862
21.4071 | -100.8614
21.4108 | -100.8587
21.4145 | -100.8548
21.4182 | -100.847
21.4219 | -100.8382
21.4255 | -100.8302
21.4292 | -100.82
21.4329 | -100.8096
21.4366 | -100.7998
21.4403 | -100.7893
21.444 | -100.778
21.4477 | -100.7651
21.4514 | -100.7512
21.455 | -100.735
21.4587 | -100.7162
21.4624 | -100.6955
21.4661 | -100.674
21.4698 | -100.6537
21.4735 | -100.634
21.4772 | -100.6131
21.4809 | -100.5907
21.4846 | -100.5684
21.4882 | -100.5436
21.4919 | -100.5199
21.4956 | -100.4979
21.4993 | -100.4738
21.503 | -100.4473
21.5067 | -100.4211
21.5104 | -100.3918
21.5141 | -100.3658
21.5177 | -100.3419
21.5214 | -100.3172
21.5251 | -100.2932
21.5288 | -100.2666
21.5325 | -100.2418
21.5362 | -100.2179
21.5399 | -100.1932
21.5436 | -100.1688
21.5473 | -100.1434
21.5509 | -100.1193
21.5546 | -100.097
21.5583 | -100.0759
21.562 | -100.0523
21.5657 | -100.031
21.5694 | -100.0119
21.5731 | -99.993
21.5768 | -99.9697
21.5805 | -99.9501
21.5841 | -99.9319
21.5878 | -99.9157
21.5915 | -99.8962
21.5952 | -99.88
21.5989 | -99.8654
21.6026 | -99.8515
21.6063 | -99.8394
21.61 | -99.8292
21.6136 | -99.8197
21.6173 | -99.8098
21.621 | -99.8028
21.6247 | -99.7969
21.6284 | -99.7918
21.6321 | -99.7902
21.6358 | -99.7889
21.6395 | -99.7896
21.6432 | -99.7899
21.6468 | -99.7926
21.6505 | -99.7963
21.6542 | -99.7994
21.6579 | -99.8043
21.6616 | -99.8125
21.6653 | -99.8221
21.669 | -99.8336
21.6727 | -99.8464
21.6763 | -99.8589
21.68 | -99.872
21.6837 | -99.8861
21.6874 | -99.9025
21.6911 | -99.9202
21.6948 | -99.9373
21.6985 | -99.9537
21.7022 | -99.9725
21.7059 | -99.992
21.7095 | -100.0128
21.7132 | -100.0356
21.7169 | -100.0563
21.7206 | -100.078
21.7243 | -100.1029
21.728 | -100.1279
21.7317 | -100.1527
21.7354 | -100.1748
21.739 | -100.1944
21.7427 | -100.2131
21.7464 | -100.2322
21.7501 | -100.2516
21.7538 | -100.2725
21.7575 | -100.2937
21.7612 | -100.315
21.7649 | -100.336
21.7686 | -100.3565
21.7722 | -100.3765
21.7759 | -100.3926
21.7796 | -100.4053
21.7833 | -100.4183
21.787 | -100.4341
21.7907 | -100.446
21.7944 | -100.4552
21.7981 | -100.4625
21.8018 | -100.4668
21.8054 | -100.472
21.8091 | -100.4779
21.8128 | -100.4817
21.8165 | -100.4773
21.8202 | -100.4828
21.8239 | -100.4836
21.8276 | -100.4808
21.8313 | -100.4777
21.8349 | -100.475
21.8386 | -100.474
21.8423 | -100.4693
21.846 | -100.4636
21.8497 | -100.4584
21.8534 | -100.4546
21.8571 | -100.4504
21.8608 | -100.4455
21.8645 | -100.4393
21.8681 | -100.4306
21.8718 | -100.4254
21.8755 | -100.4204
21.8792 | -100.4151
21.8829 | -100.411
21.8866 | -100.4129
21.8903 | -100.412
21.894 | -100.4112
21.8976 | -100.4107
21.9013 | -100.4127
21.905 | -100.4136
21.9087 | -100.4143
21.9124 | -100.4168
21.9161 | -100.42
21.9198 | -100.4245
21.9235 | -100.4328
21.9272 | -100.4434
21.9308 | -100.4512
21.9345 | -100.4614
21.9382 | -100.4765
21.9419 | -100.491
21.9456 | -100.505
21.9493 | -100.5182
21.953 | -100.5328
21.9567 | -100.5508
21.9603 | -100.5703
21.964 | -100.5922
21.9677 | -100.6156
21.9714 | -100.6384
21.9751 | -100.6617
21.9788 | -100.6837
21.9825 | -100.7015
21.9862 | -100.7216
21.9899 | -100.7451
21.9935 | -100.7671
21.9972 | -100.7891
22.0009 | -100.811
22.0046 | -100.832
22.0083 | -100.8556
22.012 | -100.8791
22.0157 | -100.9016
22.0194 | -100.9231
22.0231 | -100.9426
22.0267 | -100.9614
22.0304 | -100.9813
22.0341 | -100.9999
22.0378 | -101.0172
22.0415 | -101.0325
22.0452 | -101.0485
22.0489 | -101.0653
22.0526 | -101.0821
22.0562 | -101.0974
22.0599 | -101.1131
22.0636 | -101.125
22.0673 | -101.1362
22.071 | -101.1464
22.0747 | -101.1557
22.0784 | -101.167
22.0821 | -101.1766
22.0858 | -101.1866
22.0894 | -101.2
22.0931 | -101.2136
22.0968 | -101.2237
22.1005 | -101.2356
22.1042 | -101.2507
22.1079 | -101.2671
22.1116 | -101.2835
22.1153 | -101.3006
22.1189 | -101.3158
22.1226 | -101.3289
22.1263 | -101.3422
22.13 | -101.358
22.1337 | -101.3733
22.1374 | -101.3911
22.1411 | -101.4115
22.1448 | -101.4317
22.1485 | -101.4514
22.1521 | -101.4729
22.1558 | -101.496
22.1595 | -101.5197
22.1632 | -101.5433
22.1669 | -101.5663
22.1706 | -101.5912
22.1743 | -101.6168
22.178 | -101.6419
22.1817 | -101.6683
22.1853 | -101.6929
22.189 | -101.7161
22.1927 | -101.7405
22.1964 | -101.7661
22.2001 | -101.7896
22.2038 | -101.8143
22.2075 | -101.8417
22.2112 | -101.8685
22.2148 | -101.8982
22.2185 | -101.9282
22.2222 | -101.9572
22.2259 | -101.9848
22.2296 | -102.0111
22.2333 | -102.0365
22.237 | -102.0605
22.2407 | -102.0861
22.2444 | -102.115
22.248 | -102.1431
22.2517 | -102.167
22.2554 | -102.1921
22.2591 | -102.2176
22.2628 | -102.247
22.2665 | -102.2745
22.2702 | -102.2975
22.2739 | -102.3221
22.2775 | -102.3461
22.2812 | -102.3686
22.2849 | -102.3909
22.2886 | -102.4143
22.2923 | -102.4382
22.296 | -102.4669
22.2997 | -102.4947
22.3034 | -102.5166
22.3071 | -102.5383
22.3107 | -102.5615
22.3144 | -102.5855
22.3181 | -102.611
22.3218 | -102.6362
22.3255 | -102.6583
22.3292 | -102.6779
22.3329 | -102.6992
22.3366 | -102.7243
22.3402 | -102.7509
22.3439 | -102.7732
22.3476 | -102.7946
22.3513 | -102.8157
22.355 | -102.8385
22.3587 | -102.86
22.3624 | -102.8828
22.3661 | -102.9056
22.3698 | -102.9274
22.3734 | -102.948
22.3771 | -102.9683
22.3808 | -102.9883
22.3845 | -103.0068
22.3882 | -103.0268
22.3919 | -103.0488
22.3956 | -103.0672
22.3993 | -103.0857
22.403 | -103.1058
22.4066 | -103.1217
22.4103 | -103.1356
22.414 | -103.1507
22.4177 | -103.1639
22.4214 | -103.174
22.4251 | -103.1846
22.4288 | -103.1922
22.4325 | -103.1985
22.4361 | -103.2069
22.4398 | -103.2144
22.4435 | -103.2193
22.4472 | -103.221
22.4509 | -103.2211
22.4546 | -103.2127
22.4583 | -103.2016
22.462 | -103.1895
22.4657 | -103.1786
22.4693 | -103.1691
22.473 | -103.1587
22.4767 | -103.1452
22.4804 | -103.1283
22.4841 | -103.1113
22.4878 | -103.0942
22.4915 | -103.0754
22.4952 | -103.0543
22.4988 | -103.0329
22.5025 | -103.0118
22.5062 | -102.9883
22.5099 | -102.9602
22.5136 | -102.9307
22.5173 | -102.9021
22.521 | -102.8712
22.5247 | -102.8368
22.5284 | -102.8016
22.532 | -102.7657
22.5357 | -102.7285
22.5394 | -102.6922
22.5431 | -102.6564
22.5468 | -102.623
22.5505 | -102.5908
22.5542 | -102.5582
22.5579 | -102.5255
22.5615 | -102.4939
22.5652 | -102.4635
22.5689 | -102.435
22.5726 | -102.4063
22.5763 | -102.3795
22.58 | -102.3553
22.5837 | -102.3317
22.5874 | -102.3083
22.5911 | -102.289
22.5947 | -102.273
22.5984 | -102.2591
22.6021 | -102.2467
22.6058 | -102.2367
22.6095 | -102.2306
22.6132 | -102.2256
22.6169 | -102.2202
22.6206 | -102.2187
22.6243 | -102.2212
22.6279 | -102.2244
22.6316 | -102.2305
22.6353 | -102.2396
22.639 | -102.25
22.6427 | -102.263
22.6464 | -102.2795
22.6501 | -102.2947
22.6538 | -102.3106
22.6574 | -102.3301
22.6611 | -102.3519
22.6648 | -102.3744
22.6685 | -102.3993
22.6722 | -102.4243
22.6759 | -102.4484
22.6796 | -102.4738
22.6833 | -102.4985
22.687 | -102.5238
22.6906 | -102.5484
22.6943 | -102.5732
22.698 | -102.5973
22.7017 | -102.6221
22.7054 | -102.6457
22.7091 | -102.6658
22.7128 | -102.6867
22.7165 | -102.7081
22.7201 | -102.7279
22.7238 | -102.7462
22.7275 | -102.7635
22.7312 | -102.7802
22.7349 | -102.793
22.7386 | -102.8062
22.7423 | -102.82
22.746 | -102.8314
22.7497 | -102.8443
22.7533 | -102.8557
22.757 | -102.865
22.7607 | -102.8718
22.7644 | -102.8774
22.7681 | -102.8809
22.7718 | -102.8854
22.7755 | -102.8941
22.7792 | -102.898
22.7828 | -102.899
22.7865 | -102.8975
22.7902 | -102.8918
22.7939 | -102.8873
22.7976 | -102.883
22.8013 | -102.8772
22.805 | -102.8701
22.8087 | -102.8594
22.8124 | -102.8484
22.816 | -102.8407
22.8197 | -102.8334
22.8234 | -102.8238
22.8271 | -102.8133
22.8308 | -102.8052
22.8345 | -102.7942
22.8382 | -102.781
22.8419 | -102.7698
22.8456 | -102.7609
22.8492 | -102.7507
22.8529 | -102.7397
22.8566 | -102.7296
22.8603 | -102.7184
22.864 | -102.7049
22.8677 | -102.6905
22.8714 | -102.6753
22.8751 | -102.6587
22.8787 | -102.6395
22.8824 | -102.6239
22.8861 | -102.6092
22.8898 | -102.5932
22.8935 | -102.5755
22.8972 | -102.5587
22.9009 | -102.542
22.9046 | -102.5236
22.9083 | -102.5071
22.9119 | -102.4887
22.9156 | -102.4701
22.9193 | -102.4496
22.923 | -102.4261
22.9267 | -102.405
22.9304 | -102.3853
22.9341 | -102.3661
22.9378 | -102.3475
22.9414 | -102.3314
22.9451 | -102.3151
22.9488 | -102.2961
22.9525 | -102.2765
22.9562 | -102.2571
22.9599 | -102.2382
22.9636 | -102.2179
22.9673 | -102.1992
22.971 | -102.1828
22.9746 | -102.1684
22.9783 | -102.1541
22.982 | -102.1383
22.9857 | -102.1233
22.9894 | -102.109
22.9931 | -102.0936
22.9968 | -102.0797
23.0005 | -102.0672
23.0041 | -102.0515
23.0078 | -102.0367
23.0115 | -102.0234
23.0152 | -102.0122
23.0189 | -102.003
23.0226 | -101.9929
23.0263 | -101.9813
23.03 | -101.9732
23.0337 | -101.9671
23.0373 | -101.9601
23.041 | -101.9532
23.0447 | -101.9482
23.0484 | -101.9419
23.0521 | -101.9374
23.0558 | -101.9319
23.0595 | -101.9275
23.0632 | -101.923
23.0669 | -101.9187
23.0705 | -101.9149
23.0742 | -101.9093
23.0779 | -101.9051
23.0816 | -101.9016
23.0853 | -101.9001
23.089 | -101.8981
23.0927 | -101.8947
23.0964 | -101.8922
23.1 | -101.8908
23.1037 | -101.8897
23.1074 | -101.889
23.1111 | -101.8875
23.1148 | -101.8859
23.1185 | -101.882
23.1222 | -101.8781
23.1259 | -101.8777
23.1296 | -101.8775
23.1332 | -101.8786
23.1369 | -101.878
23.1406 | -101.8761
23.1443 | -101.8738
23.148 | -101.8728
23.1517 | -101.8727
23.1554 | -101.8721
23.1591 | -101.8721
23.1627 | -101.8734
23.1664 | -101.8741
23.1701 | -101.8746
23.1738 | -101.8761
23.1775 | -101.8781
23.1812 | -101.8795
23.1849 | -101.8813
23.1886 | -101.8841
23.1923 | -101.887
23.1959 | -101.8881
23.1996 | -101.8925
23.2033 | -101.8975
23.207 | -101.9019
23.2107 | -101.9065
23.2144 | -101.9105
23.2181 | -101.9168
23.2218 | -101.9244
23.2254 | -101.9326
23.2291 | -101.9394
23.2328 | -101.948
23.2365 | -101.9562
23.2402 | -101.9662
23.2439 | -101.9749
23.2476 | -101.9839
23.2513 | -101.9945
23.255 | -102.0025
23.2586 | -102.009
23.2623 | -102.0143
23.266 | -102.0207
23.2697 | -102.0265
23.2734 | -102.0321
23.2771 | -102.038
23.2808 | -102.0453
23.2845 | -102.0524
23.2882 | -102.0575
23.2918 | -102.0662
23.2955 | -102.0741
23.2992 | -102.08
23.3029 | -102.0873
23.3066 | -102.0935
23.3103 | -102.101
23.314 | -102.1074
23.3177 | -102.1144
23.3213 | -102.123
23.325 | -102.1311
23.3287 | -102.1392
23.3324 | -102.1496
23.3361 | -102.1579
23.3398 | -102.1654
23.3435 | -102.1695
23.3472 | -102.1749
23.3509 | -102.1802
23.3545 | -102.1872
23.3582 | -102.1925
23.3619 | -102.1972
23.3656 | -102.2035
23.3693 | -102.2098
23.373 | -102.2156
23.3767 | -102.2238
23.3804 | -102.2297
23.384 | -102.2358
23.3877 | -102.2429
23.3914 | -102.2512
23.3951 | -102.2581
23.3988 | -102.2646
23.4025 | -102.2745
23.4062 | -102.2824
23.4099 | -102.2903
23.4136 | -102.2969
23.4172 | -102.3039
23.4209 | -102.3093
23.4246 | -102.316
23.4283 | -102.3212
23.432 | -102.3251
23.4357 | -102.3306
23.4394 | -102.3366
23.4431 | -102.3392
23.4468 | -102.3421
23.4504 | -102.3478
23.4541 | -102.3525
23.4578 | -102.3565
23.4615 | -102.3605
23.4652 | -102.3635
23.4689 | -102.3678
23.4726 | -102.372
23.4763 | -102.3759
23.4799 | -102.3799
23.4836 | -102.3805
23.4873 | -102.3784
23.491 | -102.379
23.4947 | -102.3798
23.4984 | -102.3784
23.5021 | -102.3781
23.5058 | -102.3774
23.5095 | -102.3772
23.5131 | -102.3773
23.5168 | -102.3768
23.5205 | -102.377
23.5242 | -102.3755
23.5279 | -102.3746
23.5316 | -102.3734
23.5353 | -102.3721
23.539 | -102.3686
23.5426 | -102.3643
23.5463 | -102.36
23.55 | -102.3551
23.5537 | -102.3475
23.5574 | -102.3406
23.5611 | -102.3364
23.5648 | -102.3321
23.5685 | -102.3268
23.5722 | -102.3211
23.5758 | -102.3148
23.5795 | -102.3103
23.5832 | -102.3079
23.5869 | -102.3035
23.5906 | -102.2993
23.5943 | -102.2979
23.598 | -102.2943
23.6017 | -102.2885
23.6053 | -102.284
23.609 | -102.2797
23.6127 | -102.276
23.6164 | -102.2703
23.6201 | -102.2625
23.6238 | -102.2534
23.6275 | -102.2457
23.6312 | -102.24
23.6349 | -102.2322
23.6385 | -102.2264
23.6422 | -102.2194
23.6459 | -102.2134
23.6496 | -102.2066
23.6533 | -102.1996
23.657 | -102.1926
23.6607 | -102.1846
23.6644 | -102.1757
23.6681 | -102.169
23.6717 | -102.1613
23.6754 | -102.1526
23.6791 | -102.1415
23.6828 | -102.1339
23.6865 | -102.1231
23.6902 | -102.1105
23.6939 | -102.101
23.6976 | -102.0909
23.7012 | -102.0803
23.7049 | -102.0718
23.7086 | -102.0618
23.7123 | -102.0502
23.716 | -102.0403
23.7197 | -102.0296
23.7234 | -102.0175
23.7271 | -102.0062
23.7308 | -101.9954
23.7344 | -101.9843
23.7381 | -101.9748
23.7418 | -101.9643
23.7455 | -101.953
23.7492 | -101.9428
23.7529 | -101.9333
23.7566 | -101.9219
23.7603 | -101.9094
23.7639 | -101.899
23.7676 | -101.8874
23.7713 | -101.8762
23.775 | -101.8641
23.7787 | -101.8534
23.7824 | -101.8425
23.7861 | -101.8299
23.7898 | -101.8183
23.7935 | -101.806
23.7971 | -101.793
23.8008 | -101.78
23.8045 | -101.7681
23.8082 | -101.7554
23.8119 | -101.7414
23.8156 | -101.7256
23.8193 | -101.7159
23.823 | -101.7079
23.8266 | -101.697
23.8303 | -101.6852
23.834 | -101.674
23.8377 | -101.6606
23.8414 | -101.65
23.8451 | -101.6393
23.8488 | -101.629
23.8525 | -101.6187
23.8562 | -101.6086
23.8598 | -101.5992
23.8635 | -101.5908
23.8672 | -101.5816
23.8709 | -101.5716
23.8746 | -101.5633
23.8783 | -101.5558
23.882 | -101.55
23.8857 | -101.5443
23.8894 | -101.5395
23.893 | -101.5343
23.8967 | -101.5275
23.9004 | -101.5204
23.9041 | -101.5152
23.9078 | -101.511
23.9115 | -101.5049
23.9152 | -101.4987
23.9189 | -101.4916
23.9225 | -101.4838
23.9262 | -101.4784
23.9299 | -101.4738
23.9336 | -101.4701
23.9373 | -101.4675
23.941 | -101.4642
23.9447 | -101.4627
23.9484 | -101.4593
23.9521 | -101.4567
23.9557 | -101.4556
23.9594 | -101.457
23.9631 | -101.4578
23.9668 | -101.4572
23.9705 | -101.4541
23.9742 | -101.4541
23.9779 | -101.4535
23.9816 | -101.4523
23.9852 | -101.451
23.9889 | -101.4484
23.9926 | -101.4453
23.9963 | -101.4444
24 | -101.4463
24.0037 | -101.4472
24.0074 | -101.4456
24.0111 | -101.4455
24.0148 | -101.4463
24.0184 | -101.4475
24.0221 | -101.4461
24.0258 | -101.4481
24.0295 | -101.4508
24.0332 | -101.4504
24.0369 | -101.45
24.0406 | -101.4501
24.0443 | -101.4479
24.0479 | -101.4448
24.0516 | -101.4422
24.0553 | -101.4381
24.059 | -101.4338
24.0627 | -101.4312
24.0664 | -101.4284
24.0701 | -101.4258
24.0738 | -101.4232
24.0775 | -101.4207
24.0811 | -101.4177
24.0848 | -101.4173
24.0885 | -101.4142
24.0922 | -101.4107
24.0959 | -101.4107
24.0996 | -101.411
24.1033 | -101.4093
24.107 | -101.4074
24.1106 | -101.4058
24.1143 | -101.4047
24.118 | -101.4043
24.1217 | -101.4038
24.1254 | -101.4051
24.1291 | -101.407
24.1328 | -101.4078
24.1365 | -101.4071
24.1402 | -101.4074
24.1438 | -101.4079
24.1475 | -101.4102
24.1512 | -101.4144
24.1549 | -101.4194
24.1586 | -101.4213
24.1623 | -101.4241
24.166 | -101.4273
24.1697 | -101.4298
24.1734 | -101.4306
24.177 | -101.435
24.1807 | -101.4385
24.1844 | -101.4413
24.1881 | -101.4449
24.1918 | -101.448
24.1955 | -101.4484
24.1992 | -101.4524
24.2029 | -101.4571
24.2065 | -101.4584
24.2102 | -101.462
24.2139 | -101.4666
24.2176 | -101.4681
24.2213 | -101.4698
24.225 | -101.4722
24.2287 | -101.4767
24.2324 | -101.4791
24.2361 | -101.4789
24.2397 | -101.4774
24.2434 | -101.4759
24.2471 | -101.4737
24.2508 | -101.4734
24.2545 | -101.4713
24.2582 | -101.4681
24.2619 | -101.4651
24.2656 | -101.457
24.2692 | -101.4408
24.2729 | -101.4049
24.2766 | -101.3455
24.2803 | -101.2695
24.284 | -101.1845
24.2877 | -101.127
24.2914 | -101.0843
24.2951 | -101.0365
24.2988 | -100.9853
24.3024 | -100.932
24.3061 | -100.8737
24.3098 | -100.8114
24.3135 | -100.7469
24.3172 | -100.6791
24.3209 | -100.609
24.3246 | -100.5379
24.3283 | -100.4637
24.3319 | -100.3877
24.3356 | -100.3103
24.3393 | -100.2287
24.343 | -100.1436
24.3467 | -100.0587
24.3504 | -99.9737
24.3541 | -99.8876
24.3578 | -99.8019
24.3615 | -99.7154
24.3651 | -99.6301
24.3688 | -99.5451
24.3725 | -99.4612
24.3762 | -99.3796
24.3799 | -99.2993
24.3836 | -99.2196
24.3873 | -99.1421
24.391 | -99.0687
24.3947 | -98.9998
24.3983 | -98.933
24.402 | -98.8706
24.4057 | -98.811
24.4094 | -98.7545
24.4131 | -98.7012
24.4168 | -98.6525
24.4205 | -98.607
24.4242 | -98.5676
24.4278 | -98.5337
24.4315 | -98.5049
24.4352 | -98.4813
24.4389 | -98.462
24.4426 | -98.4469
24.4463 | -98.4362
24.45 | -98.4314
24.4537 | -98.4303
24.4574 | -98.4342
24.461 | -98.4439
24.4647 | -98.4592
24.4684 | -98.4795
24.4721 | -98.5037
24.4758 | -98.5651
24.4795 | -98.5962
24.4832 | -98.6322
24.4869 | -98.6719
24.4905 | -98.7155
24.4942 | -98.7622
24.4979 | -98.8116
24.5016 | -98.8633
24.5053 | -98.9171
24.509 | -98.9737
24.5127 | -99.0332
24.5164 | -99.0947
24.5201 | -99.1584
24.5237 | -99.2222
24.5274 | -99.2859
24.5311 | -99.351
24.5348 | -99.4168
24.5385 | -99.4827
24.5422 | -99.5492
24.5459 | -99.6124
24.5496 | -99.6695
24.5532 | -99.7218
24.5569 | -99.7733
24.5606 | -99.8236
24.5643 | -99.8689
24.568 | -99.9039
24.5717 | -99.9358
24.5754 | -99.9649
24.5791 | -99.9797
24.5828 | -99.9758
24.5864 | -99.9513
24.5901 | -99.9341
24.5938 | -99.9261
24.5975 | -99.9136
24.6012 | -99.8964
24.6049 | -99.8749
24.6086 | -99.8487
24.6123 | -99.8182
24.616 | -99.7848
24.6196 | -99.7566
24.6233 | -99.7339
24.627 | -99.7103
24.6307 | -99.678
24.6344 | -99.6148
24.6381 | -99.5354
24.6418 | -99.4547
24.6455 | -99.3729
24.6491 | -99.2888
24.6528 | -99.2046
24.6565 | -99.1202
24.6602 | -99.0346
24.6639 | -98.9475
24.6676 | -98.8607
24.6713 | -98.7745
24.675 | -98.6888
24.6787 | -98.6036
24.6823 | -98.5204
24.686 | -98.4398
24.6897 | -98.361
24.6934 | -98.2847
24.6971 | -98.211
24.7008 | -98.1398
24.7045 | -98.072
24.7082 | -98.0079
24.7118 | -97.9481
24.7155 | -97.8925
24.7192 | -97.841
24.7229 | -97.7939
24.7266 | -97.7514
24.7303 | -97.7148
24.734 | -97.6838
24.7377 | -97.6579
24.7414 | -97.6365
24.745 | -97.6196
24.7487 | -97.6078
24.7524 | -97.6016
24.7561 | -97.6014
24.7598 | -97.6066
24.7635 | -97.6172
24.7672 | -97.6339
24.7709 | -97.6552
24.7746 | -97.6809
24.7782 | -97.7111
24.7819 | -97.7463
24.7856 | -97.7883
24.7893 | -97.8365
24.793 | -97.8898
24.7967 | -97.9464
24.8004 | -98.0074
24.8041 | -98.0718
24.8077 | -98.1398
24.8114 | -98.2115
24.8151 | -98.2859
24.8188 | -98.3644
24.8225 | -98.4456
24.8262 | -98.53
24.8299 | -98.6165
24.8336 | -98.7053
24.8373 | -98.7955
24.8409 | -98.8866
24.8446 | -98.9751
24.8483 | -99.0638
24.852 | -99.1556
24.8557 | -99.2493
24.8594 | -99.3421
24.8631 | -99.4336
24.8668 | -99.5238
24.8704 | -99.6141
24.8741 | -99.702
24.8778 | -99.7873
24.8815 | -99.8699
24.8852 | -99.9504
24.8889 | -100.0296
24.8926 | -100.1057
24.8963 | -100.1774
24.9 | -100.2429
24.9036 | -100.308
24.9073 | -100.3693
24.911 | -100.4244
24.9147 | -100.4785
24.9184 | -100.537
24.9221 | -100.588
24.9258 | -100.6297
24.9295 | -100.6777
24.9331 | -100.7354
24.9368 | -100.8008
24.9405 | -100.8405
24.9442 | -100.8555
24.9479 | -100.8643
24.9516 | -100.8685
24.9553 | -100.8692
24.959 | -100.8659
24.9627 | -100.8604
24.9663 | -100.8559
24.97 | -100.8502
24.9737 | -100.8413
24.9774 | -100.8304
24.9811 | -100.8206
24.9848 | -100.813
24.9885 | -100.8044
24.9922 | -100.7957
24.9959 | -100.7853
24.9995 | -100.774
25.0032 | -100.7619
25.0069 | -100.7535
25.0106 | -100.7417
25.0143 | -100.7302
25.018 | -100.7207
25.0217 | -100.7116
25.0254 | -100.7017
25.029 | -100.6935
25.0327 | -100.6875
25.0364 | -100.6795
25.0401 | -100.6731
25.0438 | -100.6679
25.0475 | -100.6628
25.0512 | -100.6586
25.0549 | -100.6565
25.0586 | -100.6545
25.0622 | -100.654
25.0659 | -100.6541
25.0696 | -100.6545
25.0733 | -100.6557
25.077 | -100.6582
25.0807 | -100.6625
25.0844 | -100.6681
25.0881 | -100.6742
25.0917 | -100.6795
25.0954 | -100.6874
25.0991 | -100.6997
25.1028 | -100.7123
25.1065 | -100.725
25.1102 | -100.7395
25.1139 | -100.7541
25.1176 | -100.769
25.1213 | -100.7864
25.1249 | -100.8027
25.1286 | -100.8195
25.1323 | -100.8406
25.136 | -100.861
25.1397 | -100.8792
25.1434 | -100.8959
25.1471 | -100.9123
25.1508 | -100.9299
25.1544 | -100.9495
25.1581 | -100.9685
25.1618 | -100.9878
25.1655 | -101.0061
25.1692 | -101.0246
25.1729 | -101.0442
25.1766 | -101.0644
25.1803 | -101.0853
25.184 | -101.1041
25.1876 | -101.122
25.1913 | -101.138
25.195 | -101.1518
25.1987 | -101.1731
25.2024 | -101.1934
25.2061 | -101.2083
25.2098 | -101.2213
25.2135 | -101.2327
25.2172 | -101.2425
25.2208 | -101.2522
25.2245 | -101.2617
25.2282 | -101.2709
25.2319 | -101.2786
25.2356 | -101.2876
25.2393 | -101.2943
25.243 | -101.2974
25.2467 | -101.3013
25.2503 | -101.3076
25.254 | -101.3197
25.2577 | -101.3313
25.2614 | -101.3356
25.2651 | -101.3399
25.2688 | -101.343
25.2725 | -101.3407
25.2762 | -101.3401
25.2799 | -101.342
25.2835 | -101.3425
25.2872 | -101.3418
25.2909 | -101.3403
25.2946 | -101.3381
25.2983 | -101.3353
25.302 | -101.3328
25.3057 | -101.3301
25.3094 | -101.3269
25.313 | -101.3255
25.3167 | -101.3229
25.3204 | -101.3171
25.3241 | -101.3124
25.3278 | -101.3097
25.3315 | -101.3061
25.3352 | -101.301
25.3389 | -101.2952
25.3426 | -101.2879
25.3462 | -101.2815
25.3499 | -101.2776
25.3536 | -101.2733
25.3573 | -101.2707
25.361 | -101.2668
25.3647 | -101.2636
25.3684 | -101.2626
25.3721 | -101.2617
25.3757 | -101.259
25.3794 | -101.2553
25.3831 | -101.2526
25.3868 | -101.2499
25.3905 | -101.2472
25.3942 | -101.2445
25.3979 | -101.2398
25.4016 | -101.2364
25.4053 | -101.2324
25.4089 | -101.229
25.4126 | -101.223
25.4163 | -101.2148
25.42 | -101.2088
25.4237 | -101.2035
25.4274 | -101.1989
25.4311 | -101.1945
25.4348 | -101.1894
25.4385 | -101.1839
25.4421 | -101.1788
25.4458 | -101.1739
25.4495 | -101.1683
25.4532 | -101.1611
25.4569 | -101.1541
25.4606 | -101.1477
25.4643 | -101.1419
25.468 | -101.1349
25.4716 | -101.1281
25.4753 | -101.1209
25.479 | -101.1134
25.4827 | -101.1081
25.4864 | -101.1035
25.4901 | -101.0977
25.4938 | -101.0924
25.4975 | -101.0863
25.5012 | -101.0813
25.5048 | -101.0785
25.5085 | -101.0762
25.5122 | -101.0721
25.5159 | -101.0691
25.5196 | -101.0642
25.5233 | -101.0575
25.527 | -101.0521
25.5307 | -101.0481
25.5343 | -101.0439
25.538 | -101.0465
25.5417 | -101.0425
25.5454 | -101.039
25.5491 | -101.037
25.5528 | -101.0377
25.5565 | -101.0401
25.5602 | -101.039
25.5639 | -101.0386
25.5675 | -101.0406
25.5712 | -101.0416
25.5749 | -101.0431
25.5786 | -101.0475
25.5823 | -101.0531
25.586 | -101.0558
25.5897 | -101.0581
25.5934 | -101.0637
25.597 | -101.0665
25.6007 | -101.071
25.6044 | -101.0775
25.6081 | -101.0834
25.6118 | -101.0887
25.6155 | -101.0947
25.6192 | -101.0991
25.6229 | -101.1028
25.6266 | -101.1063
25.6302 | -101.1087
25.6339 | -101.1112
25.6376 | -101.1129
25.6413 | -101.1143
25.645 | -101.1176
25.6487 | -101.1215
25.6524 | -101.1244
25.6561 | -101.1259
25.6598 | -101.1284
25.6634 | -101.1308
25.6671 | -101.1328
25.6708 | -101.1343
25.6745 | -101.1353
25.6782 | -101.1359
25.6819 | -101.1371
25.6856 | -101.1358
25.6893 | -101.1339
25.6929 | -101.1331
25.6966 | -101.1325
25.7003 | -101.1302
25.704 | -101.1269
25.7077 | -101.1237
25.7114 | -101.1184
25.7151 | -101.1153
25.7188 | -101.1103
25.7225 | -101.105
25.7261 | -101.1
25.7298 | -101.0951
25.7335 | -101.0897
25.7372 | -101.083
25.7409 | -101.0773
25.7446 | -101.0699
25.7483 | -101.0578
25.752 | -101.0506
25.7556 | -101.0437
25.7593 | -101.0367
25.763 | -101.0316
25.7667 | -101.024
25.7704 | -101.0141
25.7741 | -101.0088
25.7778 | -101.0059
25.7815 | -101.0011
25.7852 | -100.994
25.7888 | -100.9872
25.7925 | -100.9813
25.7962 | -100.9751
25.7999 | -100.9696
25.8036 | -100.9645
25.8073 | -100.9587
25.811 | -100.9518
25.8147 | -100.9461
25.8183 | -100.9395
25.822 | -100.9318
25.8257 | -100.9223
25.8294 | -100.9147
25.8331 | -100.9076
25.8368 | -100.8987
25.8405 | -100.8921
25.8442 | -100.8858
25.8479 | -100.8808
25.8515 | -100.876
25.8552 | -100.8697
25.8589 | -100.8647
25.8626 | -100.8607
25.8663 | -100.8555
25.87 | -100.8469
25.8737 | -100.8393
25.8774 | -100.8323
25.8811 | -100.8254
25.8847 | -100.8188
25.8884 | -100.8138
25.8921 | -100.8059
25.8958 | -100.7982
25.8995 | -100.7911
25.9032 | -100.7811
25.9069 | -100.7709
25.9106 | -100.7629
25.9142 | -100.7555
25.9179 | -100.7476
25.9216 | -100.7404
25.9253 | -100.7336
25.929 | -100.7264
25.9327 | -100.7171
25.9364 | -100.7094
25.9401 | -100.7006
25.9438 | -100.6898
25.9474 | -100.6802
25.9511 | -100.6723
25.9548 | -100.6649
25.9585 | -100.6572
25.9622 | -100.6501
25.9659 | -100.6403
25.9696 | -100.63
25.9733 | -100.6223
25.9769 | -100.6146
25.9806 | -100.6056
25.9843 | -100.5964
25.988 | -100.5894
25.9917 | -100.5812
25.9954 | -100.5727
25.9991 | -100.5639
26.0028 | -100.555
26.0065 | -100.5465
26.0101 | -100.5393
26.0138 | -100.5304
26.0175 | -100.5204
26.0212 | -100.5121
26.0249 | -100.5022
26.0286 | -100.4944
26.0323 | -100.4864
26.036 | -100.4766
26.0397 | -100.4668
26.0433 | -100.4556
26.047 | -100.446
26.0507 | -100.4367
26.0544 | -100.4277
26.0581 | -100.4181
26.0618 | -100.4092
26.0655 | -100.4025
26.0692 | -100.3954
26.0728 | -100.3829
26.0765 | -100.3722
26.0802 | -100.3619
26.0839 | -100.3495
26.0876 | -100.3394
26.0913 | -100.3296
26.095 | -100.3184
26.0987 | -100.308
26.1024 | -100.2983
26.106 | -100.2899
26.1097 | -100.2816
26.1134 | -100.2737
26.1171 | -100.2693
26.1208 | -100.2605
26.1245 | -100.2528
26.1282 | -100.2457
26.1319 | -100.2367
26.1355 | -100.2282
26.1392 | -100.2215
26.1429 | -100.2123
26.1466 | -100.2037
26.1503 | -100.195
26.154 | -100.1881
26.1577 | -100.1783
26.1614 | -100.1686
26.1651 | -100.1607
26.1687 | -100.1519
26.1724 | -100.1414
26.1761 | -100.1345
26.1798 | -100.1311
26.1835 | -100.1242
26.1872 | -100.119
26.1909 | -100.1137
26.1946 | -100.109
26.1982 | -100.1041
26.2019 | -100.0989
26.2056 | -100.0942
26.2093 | -100.0879
26.213 | -100.0819
26.2167 | -100.0744
26.2204 | -100.0674
26.2241 | -100.0621
26.2278 | -100.0532
26.2314 | -100.0431
26.2351 | -100.0357
26.2388 | -100.0316
26.2425 | -100.0267
26.2462 | -100.022
26.2499 | -100.0181
26.2536 | -100.0133
26.2573 | -100.0063
26.261 | -100.0003
26.2646 | -99.9931
26.2683 | -99.9855
26.272 | -99.9797
26.2757 | -99.9737
26.2794 | -99.9657
26.2831 | -99.9539
26.2868 | -99.9397
26.2905 | -99.9266
26.2941 | -99.9152
26.2978 | -99.9029
26.3015 | -99.8936
26.3052 | -99.8829
26.3089 | -99.8713
26.3126 | -99.8617
26.3163 | -99.8496
26.32 | -99.8362
26.3237 | -99.8213
26.3273 | -99.8085
26.331 | -99.7947
26.3347 | -99.7825
26.3384 | -99.7702
26.3421 | -99.7515
26.3458 | -99.7333
26.3495 | -99.7178
26.3532 | -99.7022
26.3568 | -99.6863
26.3605 | -99.6712
26.3642 | -99.6545
26.3679 | -99.6359
26.3716 | -99.6198
26.3753 | -99.6042
26.379 | -99.5873
26.3827 | -99.5697
26.3864 | -99.5531
26.39 | -99.5375
26.3937 | -99.5219
26.3974 | -99.5018
26.4011 | -99.4805
26.4048 | -99.4623
26.4085 | -99.446
26.4122 | -99.431
26.4159 | -99.4154
26.4195 | -99.3975
26.4232 | -99.378
26.4269 | -99.3617
26.4306 | -99.3455
26.4343 | -99.3287
26.438 | -99.3129
26.4417 | -99.2991
26.4454 | -99.2856
26.4491 | -99.271
26.4527 | -99.255
26.4564 | -99.2395
26.4601 | -99.2239
26.4638 | -99.2078
26.4675 | -99.1925
26.4712 | -99.1774
26.4749 | -99.1637
26.4786 | -99.1486
26.4823 | -99.1334
26.4859 | -99.12
26.4896 | -99.1063
26.4933 | -99.0939
26.497 | -99.0813
26.5007 | -99.0686
26.5044 | -99.056
26.5081 | -99.046
26.5118 | -99.0345
26.5154 | -99.0205
26.5191 | -99.0088
26.5228 | -98.9989
26.5265 | -98.9898
26.5302 | -98.9843
26.5339 | -98.9789
26.5376 | -98.9729
26.5413 | -98.9677
26.545 | -98.9636
26.5486 | -98.9578
26.5523 | -98.9523
26.556 | -98.9486
26.5597 | -98.9427
26.5634 | -98.9399
26.5671 | -98.9371
26.5708 | -98.9324
26.5745 | -98.93
26.5781 | -98.9285
26.5818 | -98.9262
26.5855 | -98.9245
26.5892 | -98.924
26.5929 | -98.9241
26.5966 | -98.9247
26.6003 | -98.927
26.604 | -98.9328
26.6077 | -98.9353
26.6113 | -98.9333
26.615 | -98.9336
26.6187 | -98.9358
26.6224 | -98.938
26.6261 | -98.94
26.6298 | -98.9429
26.6335 | -98.9462
26.6372 | -98.9501
26.6408 | -98.9536
26.6445 | -98.9553
26.6482 | -98.9576
26.6519 | -98.9616
26.6556 | -98.9652
26.6593 | -98.97
26.663 | -98.9752
26.6667 | -98.9798
26.6704 | -98.9837
26.674 | -98.9884
26.6777 | -98.9894
26.6814 | -98.9903
26.6851 | -98.994
26.6888 | -98.993
26.6925 | -98.9932
26.6962 | -98.9955
26.6999 | -98.9977
26.7035 | -98.9978
26.7072 | -98.9997
26.7109 | -99
26.7146 | -99.0009
26.7183 | -99.0037
26.722 | -99.0042
26.7257 | -99.0035
26.7294 | -99.0028
26.7331 | -98.9997
26.7367 | -98.9936
26.7404 | -98.9917
26.7441 | -98.9909
26.7478 | -98.9861
26.7515 | -98.9817
26.7552 | -98.9773
26.7589 | -98.9746
26.7626 | -98.9718
26.7663 | -98.966
26.7699 | -98.9593
26.7736 | -98.957
26.7773 | -98.954
26.781 | -98.9483
26.7847 | -98.9438
26.7884 | -98.9385
26.7921 | -98.9308
26.7958 | -98.9261
26.7994 | -98.9233
26.8031 | -98.9174
26.8068 | -98.9123
26.8105 | -98.9069
26.8142 | -98.9013
26.8179 | -98.8956
26.8216 | -98.8898
26.8253 | -98.8823
26.829 | -98.8781
26.8326 | -98.876
26.8363 | -98.8724
26.84 | -98.8693
26.8437 | -98.8675
26.8474 | -98.8661
26.8511 | -98.864
26.8548 | -98.865
26.8585 | -98.8637
26.8621 | -98.8616
26.8658 | -98.8621
26.8695 | -98.8617
26.8732 | -98.8616
26.8769 | -98.8634
26.8806 | -98.866
26.8843 | -98.8681
26.888 | -98.8725
26.8917 | -98.8796
26.8953 | -98.8837
26.899 | -98.8877
26.9027 | -98.89
26.9064 | -98.8911
26.9101 | -98.8961
26.9138 | -98.9016
26.9175 | -98.9073
26.9212 | -98.9151
26.9248 | -98.9225
26.9285 | -98.929
26.9322 | -98.9388
26.9359 | -98.949
26.9396 | -98.9553
26.9433 | -98.9646
26.947 | -98.9759
26.9507 | -98.9858
26.9544 | -98.9945
26.958 | -99.0049
26.9617 | -99.0132
26.9654 | -99.023
26.9691 | -99.0338
26.9728 | -99.044
26.9765 | -99.0554
26.9802 | -99.0649
26.9839 | -99.0712
26.9876 | -99.0802
26.9912 | -99.0939
26.9949 | -99.1051
26.9986 | -99.1143
27.0023 | -99.1214
27.006 | -99.1254
27.0097 | -99.1313
27.0134 | -99.1383
27.0171 | -99.1396
27.0207 | -99.1433
27.0244 | -99.1505
27.0281 | -99.1568
27.0318 | -99.1607
27.0355 | -99.1633
27.0392 | -99.1634
27.0429 | -99.1661
27.0466 | -99.1675
27.0503 | -99.1664
27.0539 | -99.165
27.0576 | -99.165
27.0613 | -99.1646
27.065 | -99.1621
27.0687 | -99.1631
27.0724 | -99.1578
27.0761 | -99.1499
27.0798 | -99.1461
27.0834 | -99.14
27.0871 | -99.1307
27.0908 | -99.125
27.0945 | -99.1195
27.0982 | -99.1137
27.1019 | -99.1095
27.1056 | -99.1026
27.1093 | -99.0946
27.113 | -99.0897
27.1166 | -99.0859
27.1203 | -99.0766
27.124 | -99.069
27.1277 | -99.0618
27.1314 | -99.0518
27.1351 | -99.046
27.1388 | -99.0388
27.1425 | -99.0315
27.1461 | -99.0242
27.1498 | -99.0162
27.1535 | -99.0117
27.1572 | -99.0073
27.1609 | -99.0041
27.1646 | -99.0013
27.1683 | -98.9959
27.172 | -98.9907
27.1757 | -98.9884
27.1793 | -98.9871
27.183 | -98.9846
27.1867 | -98.984
27.1904 | -98.9842
27.1941 | -98.9854
27.1978 | -98.9869
27.2015 | -98.9881
27.2052 | -98.9916
27.2089 | -98.9974
27.2125 | -99.002
27.2162 | -99.0065
27.2199 | -99.012
27.2236 | -99.0189
27.2273 | -99.0253
27.231 | -99.0318
27.2347 | -99.0397
27.2384 | -99.0486
27.242 | -99.0604
27.2457 | -99.0726
27.2494 | -99.0829
27.2531 | -99.095
27.2568 | -99.1087
27.2605 | -99.122
27.2642 | -99.1369
27.2679 | -99.1532
27.2716 | -99.1689
27.2752 | -99.1831
27.2789 | -99.201
27.2826 | -99.2188
27.2863 | -99.2378
27.29 | -99.2585
27.2937 | -99.2768
27.2974 | -99.2963
27.3011 | -99.3175
27.3047 | -99.3366
27.3084 | -99.3556
27.3121 | -99.376
27.3158 | -99.3957
27.3195 | -99.4162
27.3232 | -99.4398
27.3269 | -99.4623
27.3306 | -99.4839
27.3343 | -99.5057
27.3379 | -99.5298
27.3416 | -99.5529
27.3453 | -99.5775
27.349 | -99.6002
27.3527 | -99.6216
27.3564 | -99.6426
27.3601 | -99.6636
27.3638 | -99.6866
27.3675 | -99.7107
27.3711 | -99.7329
27.3748 | -99.7531
27.3785 | -99.7736
27.3822 | -99.7938
27.3859 | -99.8117
27.3896 | -99.834
27.3933 | -99.8547
27.397 | -99.8731
27.4006 | -99.8939
27.4043 | -99.9142
27.408 | -99.9303
27.4117 | -99.9454
27.4154 | -99.961
27.4191 | -99.976
27.4228 | -99.9922
27.4265 | -100.005
27.4302 | -100.0166
27.4338 | -100.029
27.4375 | -100.0412
27.4412 | -100.0527
27.4449 | -100.0631
27.4486 | -100.0728
27.4523 | -100.0816
27.456 | -100.0946
27.4597 | -100.1065
27.4633 | -100.1128
27.467 | -100.1191
27.4707 | -100.128
27.4744 | -100.1328
27.4781 | -100.1357
27.4818 | -100.1398
27.4855 | -100.1426
27.4892 | -100.1455
27.4929 | -100.1487
27.4965 | -100.1498
27.5002 | -100.1505
27.5039 | -100.1556
27.5076 | -100.1575
27.5113 | -100.1589
27.515 | -100.1611
27.5187 | -100.1615
27.5224 | -100.1617
27.526 | -100.162
27.5297 | -100.1627
27.5334 | -100.1641
27.5371 | -100.1675
27.5408 | -100.1683
27.5445 | -100.1675
27.5482 | -100.1684
27.5519 | -100.1691
27.5556 | -100.1709
27.5592 | -100.1741
27.5629 | -100.1769
27.5666 | -100.1767
27.5703 | -100.1779
27.574 | -100.1792
27.5777 | -100.1802
27.5814 | -100.1811
27.5851 | -100.1843
27.5888 | -100.189
27.5924 | -100.1926
27.5961 | -100.1972
27.5998 | -100.2
27.6035 | -100.2015
27.6072 | -100.2027
27.6109 | -100.2017
27.6146 | -100.2035
27.6183 | -100.2052
27.6219 | -100.208
27.6256 | -100.2092
27.6293 | -100.2115
27.633 | -100.2155
27.6367 | -100.2192
27.6404 | -100.2226
27.6441 | -100.2243
27.6478 | -100.2258
27.6515 | -100.2251
27.6551 | -100.2247
27.6588 | -100.2247
27.6625 | -100.2277
27.6662 | -100.223
27.6699 | -100.2216
27.6736 | -100.2197
27.6773 | -100.217
27.681 | -100.2138
27.6846 | -100.2111
27.6883 | -100.2075
27.692 | -100.2016
27.6957 | -100.197
27.6994 | -100.1933
27.7031 | -100.1884
27.7068 | -100.1808
27.7105 | -100.172
27.7142 | -100.1632
27.7178 | -100.1547
27.7215 | -100.1458
27.7252 | -100.1346
27.7289 | -100.1232
27.7326 | -100.112
27.7363 | -100.1007
27.74 | -100.0885
27.7437 | -100.0764
27.7473 | -100.0662
27.751 | -100.0547
27.7547 | -100.0437
27.7584 | -100.0329
27.7621 | -100.0228
27.7658 | -100.012
27.7695 | -100
27.7732 | -99.9889
27.7769 | -99.9815
27.7805 | -99.9724
27.7842 | -99.9624
27.7879 | -99.9526
27.7916 | -99.9438
27.7953 | -99.933
27.799 | -99.9228
27.8027 | -99.9144
27.8064 | -99.9064
27.8101 | -99.8999
27.8137 | -99.8925
27.8174 | -99.8854
27.8211 | -99.8786
27.8248 | -99.8728
27.8285 | -99.8676
27.8322 | -99.8644
27.8359 | -99.8618
27.8396 | -99.8599
27.8432 | -99.8595
27.8469 | -99.8607
27.8506 | -99.8624
27.8543 | -99.8629
27.858 | -99.8605
27.8617 | -99.8584
27.8654 | -99.8598
27.8691 | -99.8643
27.8728 | -99.8694
27.8764 | -99.8753
27.8801 | -99.8821
27.8838 | -99.8901
27.8875 | -99.8978
27.8912 | -99.9064
27.8949 | -99.913
27.8986 | -99.9202
27.9023 | -99.9286
27.9059 | -99.9366
27.9096 | -99.9448
27.9133 | -99.9532
27.917 | -99.9597
27.9207 | -99.9653
27.9244 | -99.9675
27.9281 | -99.9597
27.9318 | -99.9536
27.9355 | -99.9553
27.9391 | -99.9563
27.9428 | -99.9582
27.9465 | -99.96
27.9502 | -99.9602
27.9539 | -99.96
27.9576 | -99.9608
27.9613 | -99.9611
27.965 | -99.9605
27.9686 | -99.9588
27.9723 | -99.9567
27.976 | -99.9527
27.9797 | -99.9472
27.9834 | -99.9387
27.9871 | -99.9293
27.9908 | -99.9218
27.9945 | -99.9139
27.9982 | -99.9042
28.0018 | -99.8949