## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 264 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 208
Hood | 02-92B/P | 148
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1007.7 | hPa |  | PRC PRO 01
Humidity | 55 | %/Hr
Internal temperature of calibration bench | 24.48 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1015 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 25.5 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.02 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.02 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.05 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3202 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 20/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 23/07/2018
Technical Manager | Anthony HUE | Approval |  | 24/07/2018
CEO | Robert GLEMOT | Approval |  | 24/07/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.3507 | 57.2629
0.012209 | -35.949 | 38.0804
0.018313 | -34.9033 | 27.9627
0.024417 | -34.4467 | 21.7465
0.030521 | -34.3379 | 18.0003
0.036626 | -34.196 | 14.9619
0.04273 | -34.1018 | 12.6107
0.048834 | -34.0573 | 11.1069
0.054938 | -33.9947 | 9.983
0.061043 | -33.9357 | 8.9115
0.067147 | -33.9261 | 8.0894
0.073251 | -33.9277 | 7.3296
0.079355 | -33.9464 | 6.6701
0.08546 | -33.961 | 6.0808
0.091564 | -33.9565 | 5.6448
0.097668 | -33.9386 | 5.3981
0.10377 | -33.9169 | 5.1519
0.10988 | -33.9 | 4.8774
0.11598 | -33.8913 | 4.5884
0.12209 | -33.8937 | 4.3085
0.12819 | -33.9042 | 4.0285
0.13429 | -33.9155 | 3.7432
0.1404 | -33.9187 | 3.467
0.1465 | -33.9206 | 3.2116
0.15261 | -33.9179 | 2.9934
0.15871 | -33.9117 | 2.8328
0.16482 | -33.9039 | 2.7306
0.17092 | -33.902 | 2.691
0.17702 | -33.9001 | 2.6187
0.18313 | -33.8993 | 2.5407
0.18923 | -33.899 | 2.4321
0.19534 | -33.8979 | 2.3058
0.20144 | -33.8968 | 2.1643
0.20754 | -33.8943 | 2.0441
0.21365 | -33.8936 | 1.9301
0.21975 | -33.8934 | 1.8328
0.22586 | -33.8947 | 1.7461
0.23196 | -33.8973 | 1.6744
0.23807 | -33.9015 | 1.6124
0.24417 | -33.9057 | 1.5539
0.25027 | -33.9098 | 1.4919
0.25638 | -33.9124 | 1.4299
0.26248 | -33.9144 | 1.363
0.26859 | -33.9155 | 1.2971
0.27469 | -33.916 | 1.2253
0.2808 | -33.917 | 1.1513
0.2869 | -33.9179 | 1.0802
0.293 | -33.9188 | 1.013
0.29911 | -33.9196 | 0.95296
0.30521 | -33.9206 | 0.89412
0.31132 | -33.9215 | 0.856
0.31742 | -33.9228 | 0.81351
0.32353 | -33.9233 | 0.77885
0.32963 | -33.9243 | 0.75013
0.33573 | -33.9255 | 0.72988
0.34184 | -33.927 | 0.70842
0.34794 | -33.9292 | 0.68397
0.35405 | -33.9314 | 0.65931
0.36015 | -33.9339 | 0.63331
0.36626 | -33.936 | 0.60503
0.37236 | -33.9382 | 0.56755
0.37846 | -33.9404 | 0.52289
0.38457 | -33.9413 | 0.46609
0.39067 | -33.9424 | 0.41945
0.39678 | -33.9431 | 0.36798
0.40288 | -33.9437 | 0.31457
0.40899 | -33.9436 | 0.26341
0.41509 | -33.9431 | 0.21341
0.42119 | -33.9431 | 0.17135
0.4273 | -33.9431 | 0.12609
0.4334 | -33.9431 | 0.080854
0.43951 | -33.9434 | 0.038096
0.44561 | -33.9443 | -0.00088548
0.45172 | -33.9459 | -0.032416
0.45782 | -33.9467 | -0.069081
0.46392 | -33.948 | -0.10496
0.47003 | -33.9491 | -0.1373
0.47613 | -33.9501 | -0.17131
0.48224 | -33.9507 | -0.20615
0.48834 | -33.9498 | -0.2414
0.49445 | -33.9497 | -0.27687
0.50055 | -33.9492 | -0.30919
0.50665 | -33.9485 | -0.34387
0.51276 | -33.9476 | -0.37834
0.51886 | -33.9466 | -0.41267
0.52497 | -33.9454 | -0.43767
0.53107 | -33.9446 | -0.46781
0.53717 | -33.9439 | -0.49659
0.54328 | -33.943 | -0.5202
0.54938 | -33.9428 | -0.54355
0.55549 | -33.9428 | -0.56527
0.56159 | -33.9438 | -0.57797
0.5677 | -33.944 | -0.59351
0.5738 | -33.945 | -0.6108
0.5799 | -33.9458 | -0.62677
0.58601 | -33.9468 | -0.63796
0.59211 | -33.9485 | -0.64536
0.59822 | -33.9498 | -0.65577
0.60432 | -33.9513 | -0.66869
0.61043 | -33.9526 | -0.67765
0.61653 | -33.9538 | -0.68869
0.62263 | -33.9554 | -0.69875
0.62874 | -33.9562 | -0.70976
0.63484 | -33.9571 | -0.71957
0.64095 | -33.9577 | -0.73016
0.64705 | -33.9582 | -0.74041
0.65316 | -33.9585 | -0.75334
0.65926 | -33.9586 | -0.76948
0.66536 | -33.9577 | -0.78724
0.67147 | -33.9574 | -0.80082
0.67757 | -33.9572 | -0.81547
0.68368 | -33.9565 | -0.83464
0.68978 | -33.9559 | -0.85216
0.69589 | -33.9551 | -0.87141
0.70199 | -33.9543 | -0.8945
0.70809 | -33.9536 | -0.91391
0.7142 | -33.953 | -0.93467
0.7203 | -33.9521 | -0.95423
0.72641 | -33.9516 | -0.97708
0.73251 | -33.9511 | -0.99746
0.73862 | -33.9514 | -1.0194
0.74472 | -33.9512 | -1.0398
0.75082 | -33.9511 | -1.0621
0.75693 | -33.9516 | -1.0842
0.76303 | -33.9522 | -1.1063
0.76914 | -33.9541 | -1.1268
0.77524 | -33.9551 | -1.1467
0.78135 | -33.9565 | -1.167
0.78745 | -33.9579 | -1.187
0.79355 | -33.9592 | -1.207
0.79966 | -33.961 | -1.2265
0.80576 | -33.9633 | -1.2463
0.81187 | -33.9649 | -1.2679
0.81797 | -33.9668 | -1.2893
0.82408 | -33.9685 | -1.31
0.83018 | -33.9703 | -1.3308
0.83628 | -33.9722 | -1.3495
0.84239 | -33.9733 | -1.3704
0.84849 | -33.9748 | -1.3923
0.8546 | -33.9762 | -1.4123
0.8607 | -33.9774 | -1.4334
0.86681 | -33.9785 | -1.4528
0.87291 | -33.9795 | -1.472
0.87901 | -33.98 | -1.4906
0.88512 | -33.9807 | -1.5091
0.89122 | -33.9814 | -1.5279
0.89733 | -33.9819 | -1.544
0.90343 | -33.9823 | -1.5621
0.90953 | -33.9825 | -1.577
0.91564 | -33.9826 | -1.5932
0.92174 | -33.9828 | -1.6087
0.92785 | -33.983 | -1.6238
0.93395 | -33.9831 | -1.6382
0.94006 | -33.9833 | -1.6515
0.94616 | -33.9831 | -1.6638
0.95226 | -33.9833 | -1.677
0.95837 | -33.9834 | -1.6911
0.96447 | -33.9836 | -1.7044
0.97058 | -33.9835 | -1.7181
0.97668 | -33.9834 | -1.7319
0.98279 | -33.9831 | -1.7449
0.98889 | -33.9829 | -1.7559
0.99499 | -33.9829 | -1.7697
1.0011 | -33.9829 | -1.7808
1.0072 | -33.9825 | -1.7929
1.0133 | -33.9822 | -1.8039
1.0194 | -33.9819 | -1.8172
1.0255 | -33.9816 | -1.8312
1.0316 | -33.981 | -1.8434
1.0377 | -33.9806 | -1.8573
1.0438 | -33.9801 | -1.8703
1.0499 | -33.9793 | -1.8858
1.056 | -33.9787 | -1.9003
1.0621 | -33.9781 | -1.9149
1.0682 | -33.9776 | -1.9277
1.0743 | -33.977 | -1.9434
1.0805 | -33.9765 | -1.9581
1.0866 | -33.9756 | -1.9736
1.0927 | -33.9747 | -1.9883
1.0988 | -33.9741 | -2.0035
1.1049 | -33.9734 | -2.0181
1.111 | -33.9726 | -2.0339
1.1171 | -33.9719 | -2.0487
1.1232 | -33.9714 | -2.0644
1.1293 | -33.971 | -2.0788
1.1354 | -33.9708 | -2.0942
1.1415 | -33.9704 | -2.1101
1.1476 | -33.9698 | -2.126
1.1537 | -33.9693 | -2.1411
1.1598 | -33.9694 | -2.1564
1.1659 | -33.9692 | -2.1702
1.172 | -33.969 | -2.1846
1.1781 | -33.9688 | -2.199
1.1842 | -33.9686 | -2.2136
1.1903 | -33.969 | -2.225
1.1964 | -33.9689 | -2.2388
1.2025 | -33.969 | -2.2532
1.2086 | -33.9693 | -2.2661
1.2147 | -33.9693 | -2.2791
1.2209 | -33.9697 | -2.2907
1.227 | -33.9705 | -2.3019
1.2331 | -33.9709 | -2.3144
1.2392 | -33.9713 | -2.3267
1.2453 | -33.9717 | -2.3381
1.2514 | -33.9721 | -2.3485
1.2575 | -33.9727 | -2.3599
1.2636 | -33.9736 | -2.3692
1.2697 | -33.9742 | -2.3789
1.2758 | -33.9749 | -2.3904
1.2819 | -33.9756 | -2.4007
1.288 | -33.9763 | -2.4119
1.2941 | -33.9772 | -2.4208
1.3002 | -33.9779 | -2.4293
1.3063 | -33.9784 | -2.4388
1.3124 | -33.9791 | -2.4482
1.3185 | -33.9799 | -2.4577
1.3246 | -33.9806 | -2.4672
1.3307 | -33.9815 | -2.4762
1.3368 | -33.9823 | -2.4853
1.3429 | -33.983 | -2.4946
1.349 | -33.9835 | -2.5044
1.3551 | -33.9842 | -2.5141
1.3613 | -33.9848 | -2.5227
1.3674 | -33.9855 | -2.5332
1.3735 | -33.986 | -2.5434
1.3796 | -33.9865 | -2.5529
1.3857 | -33.9868 | -2.5625
1.3918 | -33.9874 | -2.5731
1.3979 | -33.988 | -2.5827
1.404 | -33.9885 | -2.5945
1.4101 | -33.989 | -2.6053
1.4162 | -33.9894 | -2.6163
1.4223 | -33.9899 | -2.6281
1.4284 | -33.9904 | -2.6386
1.4345 | -33.9907 | -2.6505
1.4406 | -33.991 | -2.6644
1.4467 | -33.9915 | -2.6765
1.4528 | -33.992 | -2.6883
1.4589 | -33.9925 | -2.7005
1.465 | -33.9927 | -2.7119
1.4711 | -33.9933 | -2.7237
1.4772 | -33.9938 | -2.7374
1.4833 | -33.9943 | -2.7508
1.4894 | -33.9948 | -2.7635
1.4955 | -33.9954 | -2.7759
1.5016 | -33.996 | -2.7888
1.5078 | -33.9964 | -2.8031
1.5139 | -33.9968 | -2.8166
1.52 | -33.9972 | -2.8296
1.5261 | -33.9976 | -2.8427
1.5322 | -33.9979 | -2.8571
1.5383 | -33.9983 | -2.8701
1.5444 | -33.9989 | -2.8832
1.5505 | -33.9994 | -2.897
1.5566 | -33.9999 | -2.9105
1.5627 | -34.0004 | -2.9247
1.5688 | -34.001 | -2.9387
1.5749 | -34.0014 | -2.9519
1.581 | -34.0017 | -2.9657
1.5871 | -34.0021 | -2.9799
1.5932 | -34.0025 | -2.9932
1.5993 | -34.0029 | -3.0079
1.6054 | -34.0032 | -3.0216
1.6115 | -34.0037 | -3.0361
1.6176 | -34.004 | -3.0492
1.6237 | -34.0042 | -3.0634
1.6298 | -34.0045 | -3.0765
1.6359 | -34.0049 | -3.0897
1.642 | -34.0053 | -3.1023
1.6482 | -34.0055 | -3.1158
1.6543 | -34.0059 | -3.1299
1.6604 | -34.0062 | -3.1425
1.6665 | -34.0063 | -3.1554
1.6726 | -34.0065 | -3.1685
1.6787 | -34.0069 | -3.1816
1.6848 | -34.0068 | -3.1948
1.6909 | -34.0071 | -3.208
1.697 | -34.0073 | -3.2208
1.7031 | -34.0074 | -3.2339
1.7092 | -34.0074 | -3.2467
1.7153 | -34.0076 | -3.2592
1.7214 | -34.0074 | -3.2719
1.7275 | -34.0073 | -3.2851
1.7336 | -34.0073 | -3.2985
1.7397 | -34.0075 | -3.3113
1.7458 | -34.0075 | -3.324
1.7519 | -34.0075 | -3.3367
1.758 | -34.0076 | -3.3491
1.7641 | -34.0077 | -3.3612
1.7702 | -34.0077 | -3.3746
1.7763 | -34.0077 | -3.3866
1.7824 | -34.0077 | -3.3995
1.7885 | -34.0077 | -3.4124
1.7947 | -34.0079 | -3.4248
1.8008 | -34.0079 | -3.4366
1.8069 | -34.0079 | -3.4486
1.813 | -34.008 | -3.4613
1.8191 | -34.008 | -3.4744
1.8252 | -34.0081 | -3.4875
1.8313 | -34.0083 | -3.5002
1.8374 | -34.0085 | -3.5125
1.8435 | -34.0085 | -3.5248
1.8496 | -34.0086 | -3.5382
1.8557 | -34.0088 | -3.5513
1.8618 | -34.0091 | -3.5638
1.8679 | -34.0093 | -3.5762
1.874 | -34.0094 | -3.5886
1.8801 | -34.0096 | -3.6014
1.8862 | -34.0099 | -3.6137
1.8923 | -34.0101 | -3.6263
1.8984 | -34.0102 | -3.6386
1.9045 | -34.0103 | -3.6505
1.9106 | -34.0104 | -3.6631
1.9167 | -34.0105 | -3.6757
1.9228 | -34.0107 | -3.6879
1.9289 | -34.0109 | -3.7001
1.9351 | -34.0111 | -3.7112
1.9412 | -34.0113 | -3.7229
1.9473 | -34.0116 | -3.7353
1.9534 | -34.0119 | -3.7471
1.9595 | -34.0121 | -3.7592
1.9656 | -34.0124 | -3.7718
1.9717 | -34.0126 | -3.7839
1.9778 | -34.013 | -3.7957
1.9839 | -34.0132 | -3.8075
1.99 | -34.0135 | -3.8189
1.9961 | -34.0136 | -3.8314
2.0022 | -34.0139 | -3.8432
2.0083 | -34.014 | -3.8548
2.0144 | -34.0142 | -3.8665
2.0205 | -34.0146 | -3.8786
2.0266 | -34.0149 | -3.8904
2.0327 | -34.0152 | -3.9023
2.0388 | -34.0156 | -3.9142
2.0449 | -34.0159 | -3.926
2.051 | -34.0162 | -3.9377
2.0571 | -34.0165 | -3.9495
2.0632 | -34.0167 | -3.9613
2.0693 | -34.0169 | -3.9739
2.0754 | -34.0172 | -3.9856
2.0816 | -34.0176 | -3.9981
2.0877 | -34.0178 | -4.0101
2.0938 | -34.0182 | -4.0226
2.0999 | -34.0183 | -4.0347
2.106 | -34.0186 | -4.0467
2.1121 | -34.0188 | -4.0592
2.1182 | -34.019 | -4.0704
2.1243 | -34.0193 | -4.0816
2.1304 | -34.0195 | -4.0931
2.1365 | -34.0198 | -4.1053
2.1426 | -34.0201 | -4.1177
2.1487 | -34.0204 | -4.1294
2.1548 | -34.0208 | -4.1417
2.1609 | -34.0212 | -4.1535
2.167 | -34.0216 | -4.1652
2.1731 | -34.0219 | -4.1776
2.1792 | -34.0221 | -4.1891
2.1853 | -34.0225 | -4.2015
2.1914 | -34.0228 | -4.2132
2.1975 | -34.0232 | -4.2252
2.2036 | -34.0235 | -4.2377
2.2097 | -34.0237 | -4.2496
2.2158 | -34.0241 | -4.2606
2.222 | -34.0243 | -4.2716
2.2281 | -34.0246 | -4.2834
2.2342 | -34.0249 | -4.2953
2.2403 | -34.0251 | -4.3064
2.2464 | -34.0254 | -4.3177
2.2525 | -34.0256 | -4.3295
2.2586 | -34.026 | -4.3406
2.2647 | -34.0263 | -4.3515
2.2708 | -34.0266 | -4.3632
2.2769 | -34.0269 | -4.3744
2.283 | -34.0271 | -4.386
2.2891 | -34.0273 | -4.3973
2.2952 | -34.0275 | -4.4084
2.3013 | -34.0278 | -4.4197
2.3074 | -34.028 | -4.4306
2.3135 | -34.0282 | -4.4425
2.3196 | -34.0284 | -4.4532
2.3257 | -34.0286 | -4.4643
2.3318 | -34.0288 | -4.4758
2.3379 | -34.029 | -4.4871
2.344 | -34.0291 | -4.4989
2.3501 | -34.0293 | -4.51
2.3562 | -34.0295 | -4.5209
2.3623 | -34.0297 | -4.5318
2.3685 | -34.0298 | -4.5429
2.3746 | -34.0299 | -4.5546
2.3807 | -34.0301 | -4.5656
2.3868 | -34.0303 | -4.5768
2.3929 | -34.0305 | -4.5879
2.399 | -34.0306 | -4.5988
2.4051 | -34.0309 | -4.6098
2.4112 | -34.0311 | -4.6209
2.4173 | -34.0313 | -4.632
2.4234 | -34.0314 | -4.6433
2.4295 | -34.0316 | -4.6546
2.4356 | -34.0317 | -4.6659
2.4417 | -34.0319 | -4.6775
2.4478 | -34.032 | -4.6888
2.4539 | -34.0322 | -4.6995
2.46 | -34.0323 | -4.7115
2.4661 | -34.0325 | -4.723
2.4722 | -34.0326 | -4.7337
2.4783 | -34.0328 | -4.7446
2.4844 | -34.0328 | -4.7561
2.4905 | -34.033 | -4.7676
2.4966 | -34.0332 | -4.7787
2.5027 | -34.0333 | -4.7897
2.5089 | -34.0334 | -4.801
2.515 | -34.0335 | -4.8123
2.5211 | -34.0336 | -4.8235
2.5272 | -34.0337 | -4.8349
2.5333 | -34.0339 | -4.8463
2.5394 | -34.0341 | -4.8575
2.5455 | -34.0342 | -4.8693
2.5516 | -34.0344 | -4.8801
2.5577 | -34.0346 | -4.892
2.5638 | -34.0348 | -4.9042
2.5699 | -34.0348 | -4.9161
2.576 | -34.0349 | -4.9281
2.5821 | -34.0351 | -4.9398
2.5882 | -34.0352 | -4.9514
2.5943 | -34.0353 | -4.9627
2.6004 | -34.0356 | -4.9741
2.6065 | -34.0357 | -4.9849
2.6126 | -34.0359 | -4.9959
2.6187 | -34.036 | -5.0068
2.6248 | -34.0362 | -5.0181
2.6309 | -34.0364 | -5.03
2.637 | -34.0367 | -5.042
2.6431 | -34.0368 | -5.053
2.6492 | -34.037 | -5.0647
2.6554 | -34.0372 | -5.076
2.6615 | -34.0373 | -5.0875
2.6676 | -34.0376 | -5.0994
2.6737 | -34.0379 | -5.1109
2.6798 | -34.0382 | -5.1225
2.6859 | -34.0384 | -5.1342
2.692 | -34.0386 | -5.1455
2.6981 | -34.0388 | -5.157
2.7042 | -34.0391 | -5.1686
2.7103 | -34.0393 | -5.1796
2.7164 | -34.0396 | -5.1911
2.7225 | -34.0399 | -5.2029
2.7286 | -34.0402 | -5.2143
2.7347 | -34.0404 | -5.2253
2.7408 | -34.0407 | -5.2358
2.7469 | -34.0409 | -5.2474
2.753 | -34.0413 | -5.2584
2.7591 | -34.0416 | -5.2702
2.7652 | -34.0418 | -5.2821
2.7713 | -34.042 | -5.2933
2.7774 | -34.0423 | -5.3046
2.7835 | -34.0425 | -5.316
2.7896 | -34.0427 | -5.327
2.7958 | -34.043 | -5.3381
2.8019 | -34.0434 | -5.349
2.808 | -34.0437 | -5.3602
2.8141 | -34.044 | -5.3718
2.8202 | -34.0443 | -5.3831
2.8263 | -34.0447 | -5.3944
2.8324 | -34.045 | -5.4062
2.8385 | -34.0454 | -5.4175
2.8446 | -34.0458 | -5.4287
2.8507 | -34.0462 | -5.4399
2.8568 | -34.0466 | -5.4508
2.8629 | -34.0469 | -5.4618
2.869 | -34.0472 | -5.4733
2.8751 | -34.0475 | -5.4844
2.8812 | -34.048 | -5.4961
2.8873 | -34.0483 | -5.5074
2.8934 | -34.0486 | -5.5188
2.8995 | -34.049 | -5.5303
2.9056 | -34.0493 | -5.5418
2.9117 | -34.0496 | -5.5533
2.9178 | -34.05 | -5.5644
2.9239 | -34.0504 | -5.5764
2.93 | -34.0509 | -5.5882
2.9361 | -34.0512 | -5.5994
2.9423 | -34.0515 | -5.6105
2.9484 | -34.052 | -5.622
2.9545 | -34.0523 | -5.6334
2.9606 | -34.0528 | -5.6452
2.9667 | -34.0531 | -5.6567
2.9728 | -34.0535 | -5.6684
2.9789 | -34.0539 | -5.6799
2.985 | -34.0544 | -5.6915
2.9911 | -34.0547 | -5.7032
2.9972 | -34.0552 | -5.7147
3.0033 | -34.0556 | -5.7266
3.0094 | -34.056 | -5.7381
3.0155 | -34.0565 | -5.7493
3.0216 | -34.0568 | -5.761
3.0277 | -34.0572 | -5.773
3.0338 | -34.0576 | -5.7848
3.0399 | -34.058 | -5.7966
3.046 | -34.0583 | -5.8085
3.0521 | -34.0587 | -5.8199
3.0582 | -34.059 | -5.8323
3.0643 | -34.0593 | -5.8441
3.0704 | -34.0597 | -5.8563
3.0765 | -34.06 | -5.8678
3.0827 | -34.0603 | -5.8794
3.0888 | -34.0606 | -5.891
3.0949 | -34.061 | -5.9028
3.101 | -34.0614 | -5.9146
3.1071 | -34.0617 | -5.9267
3.1132 | -34.0621 | -5.9389
3.1193 | -34.0625 | -5.9507
3.1254 | -34.0629 | -5.9624
3.1315 | -34.0633 | -5.9747
3.1376 | -34.0635 | -5.9865
3.1437 | -34.0639 | -5.9984
3.1498 | -34.0643 | -6.0108
3.1559 | -34.0646 | -6.0227
3.162 | -34.0649 | -6.0351
3.1681 | -34.0651 | -6.0474
3.1742 | -34.0655 | -6.0593
3.1803 | -34.0658 | -6.0713
3.1864 | -34.0662 | -6.0829
3.1925 | -34.0665 | -6.0948
3.1986 | -34.0667 | -6.1071
3.2047 | -34.067 | -6.1196
3.2108 | -34.0673 | -6.131
3.2169 | -34.0676 | -6.1431
3.223 | -34.0679 | -6.1552
3.2292 | -34.0681 | -6.1674
3.2353 | -34.0683 | -6.1787
3.2414 | -34.0686 | -6.1904
3.2475 | -34.0688 | -6.2023
3.2536 | -34.0691 | -6.2141
3.2597 | -34.0694 | -6.2255
3.2658 | -34.0696 | -6.2368
3.2719 | -34.0698 | -6.2483
3.278 | -34.0701 | -6.2602
3.2841 | -34.0704 | -6.2717
3.2902 | -34.0706 | -6.2836
3.2963 | -34.0709 | -6.2954
3.3024 | -34.071 | -6.3069
3.3085 | -34.0713 | -6.3189
3.3146 | -34.0715 | -6.3305
3.3207 | -34.0718 | -6.3423
3.3268 | -34.072 | -6.3542
3.3329 | -34.0722 | -6.3653
3.339 | -34.0724 | -6.3765
3.3451 | -34.0726 | -6.3882
3.3512 | -34.0728 | -6.3993
3.3573 | -34.073 | -6.4109
3.3634 | -34.0732 | -6.4225
3.3696 | -34.0734 | -6.4342
3.3757 | -34.0736 | -6.4455
3.3818 | -34.0739 | -6.4568
3.3879 | -34.0741 | -6.4683
3.394 | -34.0742 | -6.4801
3.4001 | -34.0744 | -6.4913
3.4062 | -34.0745 | -6.5027
3.4123 | -34.0747 | -6.5137
3.4184 | -34.0749 | -6.5254
3.4245 | -34.0752 | -6.5367
3.4306 | -34.0753 | -6.5478
3.4367 | -34.0755 | -6.5588
3.4428 | -34.0757 | -6.5697
3.4489 | -34.0758 | -6.581
3.455 | -34.076 | -6.5924
3.4611 | -34.0762 | -6.604
3.4672 | -34.0764 | -6.6156
3.4733 | -34.0766 | -6.6271
3.4794 | -34.0768 | -6.6384
3.4855 | -34.077 | -6.6493
3.4916 | -34.0772 | -6.6601
3.4977 | -34.0774 | -6.6719
3.5038 | -34.0775 | -6.6831
3.5099 | -34.0777 | -6.6939
3.5161 | -34.0779 | -6.7047
3.5222 | -34.078 | -6.7156
3.5283 | -34.0781 | -6.7269
3.5344 | -34.0783 | -6.7382
3.5405 | -34.0785 | -6.7492
3.5466 | -34.0787 | -6.7606
3.5527 | -34.0788 | -6.7718
3.5588 | -34.079 | -6.7829
3.5649 | -34.0791 | -6.7941
3.571 | -34.0792 | -6.8052
3.5771 | -34.0794 | -6.8167
3.5832 | -34.0796 | -6.8271
3.5893 | -34.0798 | -6.8382
3.5954 | -34.08 | -6.8495
3.6015 | -34.0802 | -6.8609
3.6076 | -34.0804 | -6.8719
3.6137 | -34.0806 | -6.8825
3.6198 | -34.0809 | -6.8932
3.6259 | -34.081 | -6.9046
3.632 | -34.0812 | -6.9157
3.6381 | -34.0814 | -6.927
3.6442 | -34.0816 | -6.938
3.6503 | -34.0819 | -6.9487
3.6565 | -34.0821 | -6.9599
3.6626 | -34.0823 | -6.9709
3.6687 | -34.0825 | -6.9821
3.6748 | -34.0828 | -6.9929
3.6809 | -34.0831 | -7.0041
3.687 | -34.0833 | -7.0155
3.6931 | -34.0836 | -7.0262
3.6992 | -34.0838 | -7.0374
3.7053 | -34.084 | -7.0484
3.7114 | -34.0842 | -7.0595
3.7175 | -34.0845 | -7.0707
3.7236 | -34.0847 | -7.0813
3.7297 | -34.085 | -7.0925
3.7358 | -34.0852 | -7.1039
3.7419 | -34.0854 | -7.1152
3.748 | -34.0857 | -7.1262
3.7541 | -34.0859 | -7.1373
3.7602 | -34.0862 | -7.1483
3.7663 | -34.0864 | -7.1592
3.7724 | -34.0867 | -7.1704
3.7785 | -34.087 | -7.1813
3.7846 | -34.0872 | -7.1922
3.7907 | -34.0876 | -7.2036
3.7969 | -34.0878 | -7.2148
3.803 | -34.0881 | -7.2255
3.8091 | -34.0884 | -7.2365
3.8152 | -34.0887 | -7.2476
3.8213 | -34.0889 | -7.2589
3.8274 | -34.0892 | -7.2697
3.8335 | -34.0896 | -7.281
3.8396 | -34.0899 | -7.2921
3.8457 | -34.0902 | -7.3033
3.8518 | -34.0906 | -7.3143
3.8579 | -34.0909 | -7.3253
3.864 | -34.0912 | -7.3366
3.8701 | -34.0915 | -7.3481
3.8762 | -34.0918 | -7.359
3.8823 | -34.0921 | -7.3703
3.8884 | -34.0924 | -7.3814
3.8945 | -34.0928 | -7.3928
3.9006 | -34.0931 | -7.4042
3.9067 | -34.0934 | -7.4158
3.9128 | -34.0937 | -7.427
3.9189 | -34.094 | -7.438
3.925 | -34.0942 | -7.4493
3.9311 | -34.0945 | -7.46
3.9372 | -34.0949 | -7.4713
3.9434 | -34.0952 | -7.4825
3.9495 | -34.0956 | -7.4938
3.9556 | -34.0959 | -7.5047
3.9617 | -34.0962 | -7.5155
3.9678 | -34.0966 | -7.5264
3.9739 | -34.0968 | -7.5375
3.98 | -34.0971 | -7.5487
3.9861 | -34.0974 | -7.5599
3.9922 | -34.0977 | -7.571
3.9983 | -34.098 | -7.5818
4.0044 | -34.0983 | -7.5925
4.0105 | -34.0986 | -7.6036
4.0166 | -34.099 | -7.6148
4.0227 | -34.0993 | -7.6259
4.0288 | -34.0995 | -7.6375
4.0349 | -34.0998 | -7.6484
4.041 | -34.1001 | -7.6595
4.0471 | -34.1003 | -7.6707
4.0532 | -34.1006 | -7.6815
4.0593 | -34.1009 | -7.6927
4.0654 | -34.1012 | -7.7037
4.0715 | -34.1015 | -7.7152
4.0776 | -34.1019 | -7.7263
4.0838 | -34.1022 | -7.7378
4.0899 | -34.1025 | -7.7487
4.096 | -34.1028 | -7.7598
4.1021 | -34.103 | -7.7706
4.1082 | -34.1033 | -7.7815
4.1143 | -34.1036 | -7.7922
4.1204 | -34.1039 | -7.8031
4.1265 | -34.1042 | -7.8139
4.1326 | -34.1045 | -7.8247
4.1387 | -34.1048 | -7.8357
4.1448 | -34.1051 | -7.8468
4.1509 | -34.1053 | -7.8578
4.157 | -34.1056 | -7.8693
4.1631 | -34.1059 | -7.8805
4.1692 | -34.1061 | -7.8914
4.1753 | -34.1064 | -7.9024
4.1814 | -34.1066 | -7.9131
4.1875 | -34.107 | -7.9244
4.1936 | -34.1072 | -7.9355
4.1997 | -34.1075 | -7.9461
4.2058 | -34.1077 | -7.9574
4.2119 | -34.108 | -7.9683
4.218 | -34.1082 | -7.979
4.2241 | -34.1085 | -7.9898
4.2303 | -34.1088 | -8.0006
4.2364 | -34.109 | -8.0116
4.2425 | -34.1092 | -8.0226
4.2486 | -34.1095 | -8.0337
4.2547 | -34.1096 | -8.0445
4.2608 | -34.1099 | -8.0555
4.2669 | -34.1101 | -8.0662
4.273 | -34.1103 | -8.0773
4.2791 | -34.1106 | -8.0882
4.2852 | -34.1108 | -8.0993
4.2913 | -34.111 | -8.11
4.2974 | -34.1113 | -8.1212
4.3035 | -34.1115 | -8.1324
4.3096 | -34.1117 | -8.1435
4.3157 | -34.112 | -8.1543
4.3218 | -34.1122 | -8.1651
4.3279 | -34.1124 | -8.176
4.334 | -34.1126 | -8.1868
4.3401 | -34.1129 | -8.1973
4.3462 | -34.1131 | -8.2084
4.3523 | -34.1134 | -8.2189
4.3584 | -34.1136 | -8.2294
4.3645 | -34.1138 | -8.2401
4.3707 | -34.114 | -8.2507
4.3768 | -34.1143 | -8.2615
4.3829 | -34.1145 | -8.2727
4.389 | -34.1148 | -8.2835
4.3951 | -34.115 | -8.2941
4.4012 | -34.1152 | -8.3052
4.4073 | -34.1154 | -8.3158
4.4134 | -34.1156 | -8.3266
4.4195 | -34.1159 | -8.3374
4.4256 | -34.1162 | -8.348
4.4317 | -34.1164 | -8.3587
4.4378 | -34.1166 | -8.3692
4.4439 | -34.1168 | -8.3796
4.45 | -34.117 | -8.3904
4.4561 | -34.1172 | -8.4011
4.4622 | -34.1174 | -8.4119
4.4683 | -34.1176 | -8.4227
4.4744 | -34.1178 | -8.4337
4.4805 | -34.1181 | -8.444
4.4866 | -34.1183 | -8.4548
4.4927 | -34.1186 | -8.4657
4.4988 | -34.1188 | -8.4764
4.5049 | -34.119 | -8.4867
4.511 | -34.1193 | -8.4972
4.5172 | -34.1195 | -8.5079
4.5233 | -34.1197 | -8.519
4.5294 | -34.12 | -8.5296
4.5355 | -34.1202 | -8.5403
4.5416 | -34.1205 | -8.551
4.5477 | -34.1207 | -8.5617
4.5538 | -34.1209 | -8.572
4.5599 | -34.1212 | -8.5826
4.566 | -34.1214 | -8.5933
4.5721 | -34.1217 | -8.604
4.5782 | -34.122 | -8.615
4.5843 | -34.1222 | -8.6257
4.5904 | -34.1225 | -8.6365
4.5965 | -34.1228 | -8.647
4.6026 | -34.123 | -8.6581
4.6087 | -34.1233 | -8.6687
4.6148 | -34.1235 | -8.6797
4.6209 | -34.1237 | -8.6904
4.627 | -34.124 | -8.7012
4.6331 | -34.1242 | -8.7118
4.6392 | -34.1245 | -8.7226
4.6453 | -34.1247 | -8.7333
4.6514 | -34.125 | -8.7441
4.6576 | -34.1252 | -8.7551
4.6637 | -34.1254 | -8.7656
4.6698 | -34.1257 | -8.7761
4.6759 | -34.1259 | -8.7872
4.682 | -34.1261 | -8.7978
4.6881 | -34.1263 | -8.8085
4.6942 | -34.1266 | -8.8192
4.7003 | -34.1269 | -8.8299
4.7064 | -34.1271 | -8.8405
4.7125 | -34.1274 | -8.8516
4.7186 | -34.1276 | -8.8621
4.7247 | -34.1279 | -8.8729
4.7308 | -34.1282 | -8.8836
4.7369 | -34.1284 | -8.8944
4.743 | -34.1287 | -8.9055
4.7491 | -34.129 | -8.9166
4.7552 | -34.1292 | -8.9274
4.7613 | -34.1294 | -8.9383
4.7674 | -34.1297 | -8.9489
4.7735 | -34.1299 | -8.9601
4.7796 | -34.1302 | -8.9711
4.7857 | -34.1305 | -8.982
4.7918 | -34.1308 | -8.9926
4.7979 | -34.1311 | -9.0033
4.8041 | -34.1313 | -9.0141
4.8102 | -34.1316 | -9.0246
4.8163 | -34.1319 | -9.0354
4.8224 | -34.1321 | -9.0461
4.8285 | -34.1324 | -9.0566
4.8346 | -34.1327 | -9.0672
4.8407 | -34.1329 | -9.0781
4.8468 | -34.1333 | -9.0891
4.8529 | -34.1335 | -9.1002
4.859 | -34.1338 | -9.1113
4.8651 | -34.1341 | -9.1219
4.8712 | -34.1343 | -9.1325
4.8773 | -34.1346 | -9.1434
4.8834 | -34.1349 | -9.1542
4.8895 | -34.1351 | -9.1648
4.8956 | -34.1354 | -9.1756
4.9017 | -34.1356 | -9.1862
4.9078 | -34.1359 | -9.1974
4.9139 | -34.1362 | -9.2078
4.92 | -34.1365 | -9.2183
4.9261 | -34.1368 | -9.2288
4.9322 | -34.1371 | -9.2396
4.9383 | -34.1374 | -9.2503
4.9445 | -34.1377 | -9.2611
4.9506 | -34.1379 | -9.2717
4.9567 | -34.1382 | -9.2824
4.9628 | -34.1385 | -9.2931
4.9689 | -34.1388 | -9.3043
4.975 | -34.139 | -9.3154
4.9811 | -34.1393 | -9.3259
4.9872 | -34.1397 | -9.3365
4.9933 | -34.14 | -9.3476
4.9994 | -34.1403 | -9.3588
5.0055 | -34.1406 | -9.3696
5.0116 | -34.1409 | -9.3808
5.0177 | -34.1413 | -9.3919
5.0238 | -34.1416 | -9.403
5.0299 | -34.142 | -9.4141
5.036 | -34.1423 | -9.4249
5.0421 | -34.1426 | -9.4357
5.0482 | -34.1429 | -9.4465
5.0543 | -34.1432 | -9.4571
5.0604 | -34.1435 | -9.4679
5.0665 | -34.1439 | -9.4785
5.0726 | -34.1442 | -9.4897
5.0787 | -34.1445 | -9.5007
5.0848 | -34.1448 | -9.5117
5.091 | -34.1451 | -9.5224
5.0971 | -34.1454 | -9.5329
5.1032 | -34.1458 | -9.5441
5.1093 | -34.1461 | -9.5549
5.1154 | -34.1464 | -9.5655
5.1215 | -34.1466 | -9.5762
5.1276 | -34.147 | -9.5869
5.1337 | -34.1473 | -9.5975
5.1398 | -34.1476 | -9.6081
5.1459 | -34.148 | -9.6192
5.152 | -34.1483 | -9.6302
5.1581 | -34.1486 | -9.6412
5.1642 | -34.1489 | -9.6521
5.1703 | -34.1492 | -9.6628
5.1764 | -34.1496 | -9.6735
5.1825 | -34.1499 | -9.6841
5.1886 | -34.1502 | -9.6946
5.1947 | -34.1505 | -9.7054
5.2008 | -34.1509 | -9.7164
5.2069 | -34.1512 | -9.7274
5.213 | -34.1515 | -9.738
5.2191 | -34.1518 | -9.7489
5.2252 | -34.1521 | -9.7598
5.2314 | -34.1524 | -9.7709
5.2375 | -34.1528 | -9.7814
5.2436 | -34.1531 | -9.7923
5.2497 | -34.1534 | -9.8028
5.2558 | -34.1537 | -9.814
5.2619 | -34.1541 | -9.8251
5.268 | -34.1544 | -9.8361
5.2741 | -34.1548 | -9.847
5.2802 | -34.1551 | -9.8582
5.2863 | -34.1555 | -9.8689
5.2924 | -34.1558 | -9.88
5.2985 | -34.1562 | -9.8909
5.3046 | -34.1565 | -9.9018
5.3107 | -34.1568 | -9.9128
5.3168 | -34.1571 | -9.9237
5.3229 | -34.1575 | -9.9346
5.329 | -34.1578 | -9.9455
5.3351 | -34.1581 | -9.9561
5.3412 | -34.1584 | -9.9669
5.3473 | -34.1588 | -9.9778
5.3534 | -34.1591 | -9.9888
5.3595 | -34.1594 | -9.9997
5.3656 | -34.1598 | -10.0105
5.3717 | -34.1601 | -10.0216
5.3779 | -34.1605 | -10.0329
5.384 | -34.1608 | -10.0437
5.3901 | -34.1611 | -10.0549
5.3962 | -34.1615 | -10.0656
5.4023 | -34.1618 | -10.0767
5.4084 | -34.1622 | -10.0874
5.4145 | -34.1625 | -10.0981
5.4206 | -34.1629 | -10.109
5.4267 | -34.1632 | -10.1202
5.4328 | -34.1635 | -10.1311
5.4389 | -34.1639 | -10.1417
5.445 | -34.1642 | -10.153
5.4511 | -34.1645 | -10.1636
5.4572 | -34.1649 | -10.1746
5.4633 | -34.1652 | -10.1855
5.4694 | -34.1655 | -10.1964
5.4755 | -34.1658 | -10.2071
5.4816 | -34.1662 | -10.218
5.4877 | -34.1665 | -10.2288
5.4938 | -34.1669 | -10.2394
5.4999 | -34.1672 | -10.2501
5.506 | -34.1676 | -10.261
5.5121 | -34.1679 | -10.2717
5.5183 | -34.1682 | -10.2823
5.5244 | -34.1685 | -10.2934
5.5305 | -34.1689 | -10.3042
5.5366 | -34.1692 | -10.3148
5.5427 | -34.1695 | -10.3258
5.5488 | -34.1699 | -10.3367
5.5549 | -34.1702 | -10.3475
5.561 | -34.1705 | -10.3586
5.5671 | -34.1709 | -10.3693
5.5732 | -34.1712 | -10.3804
5.5793 | -34.1716 | -10.3915
5.5854 | -34.1719 | -10.4023
5.5915 | -34.1722 | -10.4135
5.5976 | -34.1726 | -10.4244
5.6037 | -34.1729 | -10.4353
5.6098 | -34.1732 | -10.4461
5.6159 | -34.1736 | -10.4568
5.622 | -34.1739 | -10.4677
5.6281 | -34.1742 | -10.4786
5.6342 | -34.1746 | -10.4898
5.6403 | -34.1749 | -10.5007
5.6464 | -34.1752 | -10.5117
5.6525 | -34.1755 | -10.5227
5.6586 | -34.1759 | -10.5332
5.6648 | -34.1762 | -10.5441
5.6709 | -34.1766 | -10.5552
5.677 | -34.1769 | -10.5661
5.6831 | -34.1773 | -10.577
5.6892 | -34.1776 | -10.5881
5.6953 | -34.1779 | -10.5989
5.7014 | -34.1783 | -10.6098
5.7075 | -34.1785 | -10.6207
5.7136 | -34.1789 | -10.6315
5.7197 | -34.1792 | -10.6424
5.7258 | -34.1795 | -10.653
5.7319 | -34.1798 | -10.6639
5.738 | -34.1802 | -10.6747
5.7441 | -34.1805 | -10.6855
5.7502 | -34.1809 | -10.6962
5.7563 | -34.1812 | -10.707
5.7624 | -34.1816 | -10.7181
5.7685 | -34.1819 | -10.729
5.7746 | -34.1822 | -10.7401
5.7807 | -34.1825 | -10.751
5.7868 | -34.1828 | -10.762
5.7929 | -34.1831 | -10.7728
5.799 | -34.1835 | -10.7835
5.8052 | -34.1838 | -10.7942
5.8113 | -34.1841 | -10.8049
5.8174 | -34.1844 | -10.8156
5.8235 | -34.1848 | -10.8264
5.8296 | -34.1851 | -10.8374
5.8357 | -34.1854 | -10.8481
5.8418 | -34.1857 | -10.8591
5.8479 | -34.186 | -10.87
5.854 | -34.1863 | -10.881
5.8601 | -34.1867 | -10.8917
5.8662 | -34.187 | -10.9027
5.8723 | -34.1873 | -10.9134
5.8784 | -34.1876 | -10.9241
5.8845 | -34.188 | -10.9348
5.8906 | -34.1883 | -10.9457
5.8967 | -34.1886 | -10.9564
5.9028 | -34.189 | -10.9671
5.9089 | -34.1894 | -10.9777
5.915 | -34.1897 | -10.9883
5.9211 | -34.19 | -10.9991
5.9272 | -34.1903 | -11.0096
5.9333 | -34.1907 | -11.0203
5.9394 | -34.191 | -11.031
5.9455 | -34.1914 | -11.0416
5.9517 | -34.1917 | -11.0523
5.9578 | -34.1921 | -11.0632
5.9639 | -34.1924 | -11.0736
5.97 | -34.1928 | -11.0845
5.9761 | -34.1932 | -11.0951
5.9822 | -34.1935 | -11.1061
5.9883 | -34.1938 | -11.1168
5.9944 | -34.1942 | -11.1277
6.0005 | -34.1945 | -11.1386
6.0066 | -34.1949 | -11.1492
6.0127 | -34.1952 | -11.1597
6.0188 | -34.1956 | -11.1707
6.0249 | -34.1959 | -11.1815
6.031 | -34.1962 | -11.192
6.0371 | -34.1966 | -11.2028
6.0432 | -34.1969 | -11.2135
6.0493 | -34.1972 | -11.2241
6.0554 | -34.1976 | -11.2347
6.0615 | -34.198 | -11.2454
6.0676 | -34.1983 | -11.2563
6.0737 | -34.1987 | -11.2672
6.0798 | -34.199 | -11.2779
6.0859 | -34.1994 | -11.2886
6.0921 | -34.1997 | -11.2991
6.0982 | -34.2 | -11.3098
6.1043 | -34.2004 | -11.3204
6.1104 | -34.2008 | -11.3313
6.1165 | -34.2011 | -11.3417
6.1226 | -34.2015 | -11.3523
6.1287 | -34.2018 | -11.3629
6.1348 | -34.2022 | -11.3736
6.1409 | -34.2025 | -11.3843
6.147 | -34.2029 | -11.3949
6.1531 | -34.2032 | -11.4056
6.1592 | -34.2035 | -11.4161
6.1653 | -34.2039 | -11.4267
6.1714 | -34.2042 | -11.4375
6.1775 | -34.2046 | -11.4483
6.1836 | -34.2049 | -11.4589
6.1897 | -34.2053 | -11.4696
6.1958 | -34.2056 | -11.4806
6.2019 | -34.206 | -11.4913
6.208 | -34.2064 | -11.502
6.2141 | -34.2067 | -11.5127
6.2202 | -34.2071 | -11.5234
6.2263 | -34.2075 | -11.5341
6.2325 | -34.2078 | -11.5445
6.2386 | -34.2081 | -11.5549
6.2447 | -34.2085 | -11.5658
6.2508 | -34.2088 | -11.5765
6.2569 | -34.2092 | -11.5871
6.263 | -34.2096 | -11.5977
6.2691 | -34.2099 | -11.6083
6.2752 | -34.2102 | -11.6191
6.2813 | -34.2106 | -11.6297
6.2874 | -34.2109 | -11.6402
6.2935 | -34.2113 | -11.6508
6.2996 | -34.2116 | -11.6614
6.3057 | -34.2119 | -11.6722
6.3118 | -34.2123 | -11.6829
6.3179 | -34.2126 | -11.6935
6.324 | -34.213 | -11.7042
6.3301 | -34.2133 | -11.7149
6.3362 | -34.2137 | -11.7254
6.3423 | -34.2141 | -11.7362
6.3484 | -34.2144 | -11.747
6.3545 | -34.2147 | -11.7574
6.3606 | -34.2151 | -11.7677
6.3667 | -34.2154 | -11.7783
6.3728 | -34.2158 | -11.7889
6.379 | -34.2161 | -11.7996
6.3851 | -34.2165 | -11.81
6.3912 | -34.2168 | -11.8206
6.3973 | -34.2171 | -11.8313
6.4034 | -34.2175 | -11.8417
6.4095 | -34.2178 | -11.8524
6.4156 | -34.2182 | -11.863
6.4217 | -34.2185 | -11.8736
6.4278 | -34.2188 | -11.8842
6.4339 | -34.2192 | -11.8948
6.44 | -34.2196 | -11.9053
6.4461 | -34.2199 | -11.9161
6.4522 | -34.2203 | -11.9265
6.4583 | -34.2206 | -11.9372
6.4644 | -34.2209 | -11.9479
6.4705 | -34.2213 | -11.9582
6.4766 | -34.2216 | -11.9689
6.4827 | -34.222 | -11.9795
6.4888 | -34.2223 | -11.9899
6.4949 | -34.2226 | -12.0003
6.501 | -34.223 | -12.011
6.5071 | -34.2233 | -12.0214
6.5132 | -34.2237 | -12.032
6.5194 | -34.224 | -12.0427
6.5255 | -34.2244 | -12.053
6.5316 | -34.2247 | -12.0638
6.5377 | -34.2251 | -12.0743
6.5438 | -34.2254 | -12.0848
6.5499 | -34.2257 | -12.0953
6.556 | -34.2261 | -12.1058
6.5621 | -34.2264 | -12.1166
6.5682 | -34.2268 | -12.1272
6.5743 | -34.2272 | -12.1378
6.5804 | -34.2275 | -12.1484
6.5865 | -34.2278 | -12.1591
6.5926 | -34.2282 | -12.1695
6.5987 | -34.2286 | -12.1799
6.6048 | -34.2289 | -12.1905
6.6109 | -34.2293 | -12.2012
6.617 | -34.2296 | -12.2119
6.6231 | -34.2299 | -12.2224
6.6292 | -34.2303 | -12.233
6.6353 | -34.2306 | -12.2438
6.6414 | -34.231 | -12.2543
6.6475 | -34.2314 | -12.2649
6.6536 | -34.2317 | -12.2755
6.6597 | -34.232 | -12.2858
6.6659 | -34.2324 | -12.2964
6.672 | -34.2328 | -12.3068
6.6781 | -34.2331 | -12.3175
6.6842 | -34.2335 | -12.3279
6.6903 | -34.2339 | -12.3383
6.6964 | -34.2342 | -12.3489
6.7025 | -34.2345 | -12.3593
6.7086 | -34.2349 | -12.3698
6.7147 | -34.2352 | -12.3805
6.7208 | -34.2356 | -12.3911
6.7269 | -34.2359 | -12.4016
6.733 | -34.2362 | -12.4119
6.7391 | -34.2366 | -12.4225
6.7452 | -34.237 | -12.4329
6.7513 | -34.2373 | -12.4432
6.7574 | -34.2377 | -12.4538
6.7635 | -34.238 | -12.4641
6.7696 | -34.2384 | -12.4746
6.7757 | -34.2388 | -12.4851
6.7818 | -34.2392 | -12.4957
6.7879 | -34.2395 | -12.5061
6.794 | -34.2399 | -12.5164
6.8001 | -34.2402 | -12.5269
6.8063 | -34.2406 | -12.5375
6.8124 | -34.2409 | -12.5481
6.8185 | -34.2412 | -12.5589
6.8246 | -34.2416 | -12.5693
6.8307 | -34.2419 | -12.5797
6.8368 | -34.2423 | -12.5904
6.8429 | -34.2426 | -12.601
6.849 | -34.243 | -12.6115
6.8551 | -34.2434 | -12.6219
6.8612 | -34.2437 | -12.6325
6.8673 | -34.2441 | -12.6429
6.8734 | -34.2444 | -12.6537
6.8795 | -34.2448 | -12.6641
6.8856 | -34.2452 | -12.6747
6.8917 | -34.2455 | -12.6852
6.8978 | -34.2458 | -12.6955
6.9039 | -34.2462 | -12.7057
6.91 | -34.2466 | -12.7163
6.9161 | -34.247 | -12.7266
6.9222 | -34.2473 | -12.737
6.9283 | -34.2477 | -12.7473
6.9344 | -34.2481 | -12.7578
6.9405 | -34.2484 | -12.7682
6.9466 | -34.2488 | -12.7789
6.9528 | -34.2492 | -12.7895
6.9589 | -34.2496 | -12.8001
6.965 | -34.2499 | -12.8107
6.9711 | -34.2503 | -12.8213
6.9772 | -34.2507 | -12.8318
6.9833 | -34.2511 | -12.8422
6.9894 | -34.2514 | -12.8529
6.9955 | -34.2518 | -12.8632
7.0016 | -34.2522 | -12.8737
7.0077 | -34.2526 | -12.8841
7.0138 | -34.2529 | -12.8945
7.0199 | -34.2533 | -12.9049
7.026 | -34.2536 | -12.9154
7.0321 | -34.254 | -12.9257
7.0382 | -34.2544 | -12.9359
7.0443 | -34.2548 | -12.9465
7.0504 | -34.2551 | -12.957
7.0565 | -34.2555 | -12.9675
7.0626 | -34.2559 | -12.9777
7.0687 | -34.2562 | -12.9878
7.0748 | -34.2566 | -12.998
7.0809 | -34.257 | -13.0084
7.087 | -34.2574 | -13.019
7.0932 | -34.2577 | -13.0294
7.0993 | -34.2581 | -13.0399
7.1054 | -34.2585 | -13.0502
7.1115 | -34.2588 | -13.0604
7.1176 | -34.2592 | -13.0708
7.1237 | -34.2595 | -13.0811
7.1298 | -34.2599 | -13.0915
7.1359 | -34.2603 | -13.1019
7.142 | -34.2606 | -13.1121
7.1481 | -34.261 | -13.1223
7.1542 | -34.2614 | -13.1328
7.1603 | -34.2617 | -13.1432
7.1664 | -34.2621 | -13.1535
7.1725 | -34.2625 | -13.1639
7.1786 | -34.2629 | -13.1743
7.1847 | -34.2633 | -13.1848
7.1908 | -34.2636 | -13.1953
7.1969 | -34.264 | -13.2058
7.203 | -34.2644 | -13.216
7.2091 | -34.2647 | -13.2263
7.2152 | -34.2651 | -13.2367
7.2213 | -34.2655 | -13.2468
7.2274 | -34.2659 | -13.2572
7.2335 | -34.2662 | -13.2675
7.2397 | -34.2666 | -13.2778
7.2458 | -34.267 | -13.2879
7.2519 | -34.2674 | -13.2982
7.258 | -34.2678 | -13.3088
7.2641 | -34.2681 | -13.3191
7.2702 | -34.2685 | -13.3297
7.2763 | -34.2689 | -13.3401
7.2824 | -34.2693 | -13.3505
7.2885 | -34.2697 | -13.3609
7.2946 | -34.27 | -13.3712
7.3007 | -34.2704 | -13.3816
7.3068 | -34.2707 | -13.3922
7.3129 | -34.2711 | -13.4026
7.319 | -34.2715 | -13.4129
7.3251 | -34.2719 | -13.4231
7.3312 | -34.2723 | -13.4334
7.3373 | -34.2727 | -13.4436
7.3434 | -34.2731 | -13.4538
7.3495 | -34.2735 | -13.464
7.3556 | -34.2739 | -13.4744
7.3617 | -34.2743 | -13.485
7.3678 | -34.2746 | -13.4952
7.3739 | -34.275 | -13.5057
7.3801 | -34.2754 | -13.5163
7.3862 | -34.2758 | -13.5266
7.3923 | -34.2761 | -13.5371
7.3984 | -34.2765 | -13.5472
7.4045 | -34.2769 | -13.5575
7.4106 | -34.2773 | -13.5677
7.4167 | -34.2777 | -13.5782
7.4228 | -34.2781 | -13.5886
7.4289 | -34.2785 | -13.5989
7.435 | -34.2789 | -13.6089
7.4411 | -34.2792 | -13.6193
7.4472 | -34.2796 | -13.6297
7.4533 | -34.28 | -13.6397
7.4594 | -34.2804 | -13.6503
7.4655 | -34.2808 | -13.6605
7.4716 | -34.2812 | -13.671
7.4777 | -34.2815 | -13.6813
7.4838 | -34.2819 | -13.6914
7.4899 | -34.2823 | -13.7016
7.496 | -34.2827 | -13.7117
7.5021 | -34.2831 | -13.722
7.5082 | -34.2835 | -13.7323
7.5143 | -34.2838 | -13.7426
7.5204 | -34.2842 | -13.7531
7.5266 | -34.2846 | -13.7632
7.5327 | -34.285 | -13.7734
7.5388 | -34.2854 | -13.7839
7.5449 | -34.2858 | -13.7944
7.551 | -34.2862 | -13.8046
7.5571 | -34.2865 | -13.815
7.5632 | -34.2869 | -13.8253
7.5693 | -34.2874 | -13.8355
7.5754 | -34.2877 | -13.8456
7.5815 | -34.2881 | -13.8559
7.5876 | -34.2885 | -13.8664
7.5937 | -34.2889 | -13.8768
7.5998 | -34.2893 | -13.8871
7.6059 | -34.2897 | -13.8976
7.612 | -34.2901 | -13.9079
7.6181 | -34.2905 | -13.9181
7.6242 | -34.2909 | -13.9286
7.6303 | -34.2913 | -13.939
7.6364 | -34.2917 | -13.9494
7.6425 | -34.292 | -13.9598
7.6486 | -34.2924 | -13.9701
7.6547 | -34.2928 | -13.9803
7.6608 | -34.2932 | -13.9906
7.667 | -34.2936 | -14.0009
7.6731 | -34.294 | -14.0113
7.6792 | -34.2944 | -14.0218
7.6853 | -34.2948 | -14.0322
7.6914 | -34.2951 | -14.0426
7.6975 | -34.2955 | -14.053
7.7036 | -34.2959 | -14.0633
7.7097 | -34.2963 | -14.0738
7.7158 | -34.2968 | -14.0841
7.7219 | -34.2971 | -14.0943
7.728 | -34.2975 | -14.1047
7.7341 | -34.2979 | -14.115
7.7402 | -34.2983 | -14.1254
7.7463 | -34.2987 | -14.1358
7.7524 | -34.2991 | -14.1463
7.7585 | -34.2995 | -14.1567
7.7646 | -34.2999 | -14.167
7.7707 | -34.3003 | -14.1774
7.7768 | -34.3006 | -14.1879
7.7829 | -34.301 | -14.1984
7.789 | -34.3014 | -14.209
7.7951 | -34.3018 | -14.2193
7.8012 | -34.3022 | -14.2297
7.8073 | -34.3026 | -14.2402
7.8135 | -34.303 | -14.2505
7.8196 | -34.3034 | -14.261
7.8257 | -34.3038 | -14.2715
7.8318 | -34.3042 | -14.2819
7.8379 | -34.3046 | -14.2922
7.844 | -34.305 | -14.3025
7.8501 | -34.3054 | -14.313
7.8562 | -34.3058 | -14.3236
7.8623 | -34.3062 | -14.3339
7.8684 | -34.3066 | -14.3443
7.8745 | -34.3069 | -14.3547
7.8806 | -34.3073 | -14.3652
7.8867 | -34.3077 | -14.3754
7.8928 | -34.3082 | -14.3855
7.8989 | -34.3086 | -14.3959
7.905 | -34.3089 | -14.4064
7.9111 | -34.3093 | -14.4168
7.9172 | -34.3097 | -14.4271
7.9233 | -34.3101 | -14.4375
7.9294 | -34.3105 | -14.448
7.9355 | -34.3109 | -14.4582
7.9416 | -34.3112 | -14.4686
7.9477 | -34.3116 | -14.4788
7.9539 | -34.312 | -14.489
7.96 | -34.3124 | -14.4995
7.9661 | -34.3128 | -14.5099
7.9722 | -34.3132 | -14.5201
7.9783 | -34.3136 | -14.5307
7.9844 | -34.314 | -14.5409
7.9905 | -34.3144 | -14.5515
7.9966 | -34.3148 | -14.562
8.0027 | -34.3152 | -14.5723
8.0088 | -34.3156 | -14.5827
8.0149 | -34.3159 | -14.5931
8.021 | -34.3163 | -14.6035
8.0271 | -34.3167 | -14.6141
8.0332 | -34.3171 | -14.6243
8.0393 | -34.3176 | -14.6348
8.0454 | -34.3179 | -14.6454
8.0515 | -34.3184 | -14.6558
8.0576 | -34.3188 | -14.6664
8.0637 | -34.3191 | -14.6766
8.0698 | -34.3195 | -14.6869
8.0759 | -34.3199 | -14.6972
8.082 | -34.3203 | -14.7075
8.0881 | -34.3207 | -14.7179
8.0942 | -34.3211 | -14.7283
8.1004 | -34.3215 | -14.7388
8.1065 | -34.3219 | -14.749
8.1126 | -34.3223 | -14.7595
8.1187 | -34.3227 | -14.77
8.1248 | -34.3231 | -14.7801
8.1309 | -34.3235 | -14.7904
8.137 | -34.3239 | -14.8005
8.1431 | -34.3243 | -14.8111
8.1492 | -34.3247 | -14.8215
8.1553 | -34.325 | -14.8319
8.1614 | -34.3254 | -14.8424
8.1675 | -34.3258 | -14.8528
8.1736 | -34.3262 | -14.8633
8.1797 | -34.3266 | -14.8737
8.1858 | -34.327 | -14.8841
8.1919 | -34.3274 | -14.8945
8.198 | -34.3278 | -14.9047
8.2041 | -34.3282 | -14.9152
8.2102 | -34.3286 | -14.9255
8.2163 | -34.329 | -14.9358
8.2224 | -34.3294 | -14.9461
8.2285 | -34.3298 | -14.9562
8.2346 | -34.3301 | -14.9666
8.2408 | -34.3306 | -14.9768
8.2469 | -34.331 | -14.9872
8.253 | -34.3314 | -14.9975
8.2591 | -34.3318 | -15.0079
8.2652 | -34.3322 | -15.0183
8.2713 | -34.3326 | -15.0288
8.2774 | -34.333 | -15.0392
8.2835 | -34.3334 | -15.0498
8.2896 | -34.3338 | -15.0599
8.2957 | -34.3342 | -15.0704
8.3018 | -34.3346 | -15.0807
8.3079 | -34.335 | -15.091
8.314 | -34.3354 | -15.1014
8.3201 | -34.3358 | -15.1117
8.3262 | -34.3362 | -15.1221
8.3323 | -34.3366 | -15.1324
8.3384 | -34.337 | -15.1428
8.3445 | -34.3374 | -15.1531
8.3506 | -34.3378 | -15.1635
8.3567 | -34.3382 | -15.1739
8.3628 | -34.3386 | -15.1843
8.3689 | -34.339 | -15.1945
8.375 | -34.3394 | -15.2049
8.3812 | -34.3398 | -15.2151
8.3873 | -34.3402 | -15.2254
8.3934 | -34.3406 | -15.2358
8.3995 | -34.3409 | -15.246
8.4056 | -34.3413 | -15.2564
8.4117 | -34.3417 | -15.2668
8.4178 | -34.3421 | -15.2771
8.4239 | -34.3425 | -15.2873
8.43 | -34.3429 | -15.2977
8.4361 | -34.3433 | -15.3081
8.4422 | -34.3437 | -15.3184
8.4483 | -34.3441 | -15.3286
8.4544 | -34.3445 | -15.339
8.4605 | -34.3449 | -15.3492
8.4666 | -34.3453 | -15.3597
8.4727 | -34.3457 | -15.37
8.4788 | -34.3461 | -15.3803
8.4849 | -34.3465 | -15.3906
8.491 | -34.3469 | -15.4009
8.4971 | -34.3473 | -15.4111
8.5032 | -34.3477 | -15.4213
8.5093 | -34.3481 | -15.4315
8.5154 | -34.3485 | -15.4419
8.5215 | -34.3489 | -15.452
8.5277 | -34.3493 | -15.4624
8.5338 | -34.3497 | -15.4726
8.5399 | -34.3501 | -15.4829
8.546 | -34.3505 | -15.4931
8.5521 | -34.3509 | -15.5034
8.5582 | -34.3513 | -15.5137
8.5643 | -34.3517 | -15.5237
8.5704 | -34.3521 | -15.5338
8.5765 | -34.3525 | -15.5442
8.5826 | -34.3529 | -15.5542
8.5887 | -34.3533 | -15.5644
8.5948 | -34.3538 | -15.5747
8.6009 | -34.3542 | -15.5848
8.607 | -34.3546 | -15.5952
8.6131 | -34.355 | -15.6055
8.6192 | -34.3554 | -15.6156
8.6253 | -34.3558 | -15.6258
8.6314 | -34.3562 | -15.636
8.6375 | -34.3566 | -15.6464
8.6436 | -34.357 | -15.6566
8.6497 | -34.3574 | -15.6668
8.6558 | -34.3578 | -15.677
8.6619 | -34.3582 | -15.6873
8.6681 | -34.3586 | -15.6977
8.6742 | -34.359 | -15.7079
8.6803 | -34.3594 | -15.7181
8.6864 | -34.3599 | -15.7281
8.6925 | -34.3603 | -15.7382
8.6986 | -34.3607 | -15.7485
8.7047 | -34.361 | -15.7589
8.7108 | -34.3614 | -15.7691
8.7169 | -34.3619 | -15.7794
8.723 | -34.3623 | -15.7895
8.7291 | -34.3627 | -15.7996
8.7352 | -34.3631 | -15.8099
8.7413 | -34.3635 | -15.8202
8.7474 | -34.3639 | -15.8303
8.7535 | -34.3643 | -15.8406
8.7596 | -34.3647 | -15.8508
8.7657 | -34.3651 | -15.8612
8.7718 | -34.3655 | -15.8715
8.7779 | -34.3659 | -15.8818
8.784 | -34.3663 | -15.8921
8.7901 | -34.3668 | -15.9023
8.7962 | -34.3672 | -15.9123
8.8023 | -34.3676 | -15.9225
8.8084 | -34.368 | -15.9326
8.8146 | -34.3684 | -15.9428
8.8207 | -34.3688 | -15.953
8.8268 | -34.3692 | -15.963
8.8329 | -34.3696 | -15.9732
8.839 | -34.3701 | -15.9834
8.8451 | -34.3705 | -15.9935
8.8512 | -34.3709 | -16.0036
8.8573 | -34.3713 | -16.0138
8.8634 | -34.3717 | -16.0238
8.8695 | -34.3721 | -16.0342
8.8756 | -34.3725 | -16.0445
8.8817 | -34.373 | -16.0545
8.8878 | -34.3734 | -16.0649
8.8939 | -34.3738 | -16.0751
8.9 | -34.3742 | -16.0852
8.9061 | -34.3746 | -16.0953
8.9122 | -34.375 | -16.1055
8.9183 | -34.3755 | -16.1157
8.9244 | -34.3759 | -16.1261
8.9305 | -34.3764 | -16.136
8.9366 | -34.3768 | -16.1462
8.9427 | -34.3772 | -16.1564
8.9488 | -34.3776 | -16.1666
8.955 | -34.378 | -16.1766
8.9611 | -34.3784 | -16.187
8.9672 | -34.3789 | -16.1972
8.9733 | -34.3793 | -16.2072
8.9794 | -34.3798 | -16.2172
8.9855 | -34.3802 | -16.2275
8.9916 | -34.3806 | -16.2378
8.9977 | -34.381 | -16.2479
9.0038 | -34.3815 | -16.2579
9.0099 | -34.3819 | -16.2681
9.016 | -34.3823 | -16.2783
9.0221 | -34.3828 | -16.2883
9.0282 | -34.3832 | -16.2983
9.0343 | -34.3836 | -16.3084
9.0404 | -34.3841 | -16.3186
9.0465 | -34.3845 | -16.3288
9.0526 | -34.385 | -16.3389
9.0587 | -34.3854 | -16.3491
9.0648 | -34.3859 | -16.3591
9.0709 | -34.3863 | -16.3692
9.077 | -34.3867 | -16.3794
9.0831 | -34.3871 | -16.3896
9.0892 | -34.3876 | -16.3998
9.0953 | -34.388 | -16.4101
9.1015 | -34.3884 | -16.4202
9.1076 | -34.3889 | -16.4302
9.1137 | -34.3893 | -16.4404
9.1198 | -34.3898 | -16.4504
9.1259 | -34.3902 | -16.4605
9.132 | -34.3906 | -16.4708
9.1381 | -34.3911 | -16.4809
9.1442 | -34.3916 | -16.4907
9.1503 | -34.392 | -16.5009
9.1564 | -34.3924 | -16.511
9.1625 | -34.3928 | -16.5212
9.1686 | -34.3933 | -16.5313
9.1747 | -34.3937 | -16.5414
9.1808 | -34.3942 | -16.5515
9.1869 | -34.3947 | -16.5616
9.193 | -34.3951 | -16.5717
9.1991 | -34.3955 | -16.5819
9.2052 | -34.396 | -16.592
9.2113 | -34.3964 | -16.6022
9.2174 | -34.3969 | -16.6123
9.2235 | -34.3973 | -16.6224
9.2296 | -34.3978 | -16.6325
9.2357 | -34.3982 | -16.6428
9.2419 | -34.3987 | -16.6528
9.248 | -34.3992 | -16.6629
9.2541 | -34.3996 | -16.673
9.2602 | -34.4001 | -16.6833
9.2663 | -34.4005 | -16.6931
9.2724 | -34.401 | -16.7032
9.2785 | -34.4014 | -16.7133
9.2846 | -34.4019 | -16.7233
9.2907 | -34.4024 | -16.7334
9.2968 | -34.4028 | -16.7437
9.3029 | -34.4032 | -16.7539
9.309 | -34.4037 | -16.764
9.3151 | -34.4041 | -16.7739
9.3212 | -34.4046 | -16.7839
9.3273 | -34.4051 | -16.794
9.3334 | -34.4055 | -16.8039
9.3395 | -34.406 | -16.8139
9.3456 | -34.4064 | -16.8238
9.3517 | -34.4068 | -16.834
9.3578 | -34.4073 | -16.844
9.3639 | -34.4078 | -16.854
9.37 | -34.4082 | -16.864
9.3761 | -34.4087 | -16.8743
9.3822 | -34.4092 | -16.8844
9.3884 | -34.4096 | -16.8943
9.3945 | -34.4101 | -16.9044
9.4006 | -34.4106 | -16.9145
9.4067 | -34.411 | -16.9246
9.4128 | -34.4115 | -16.9347
9.4189 | -34.4119 | -16.945
9.425 | -34.4124 | -16.9549
9.4311 | -34.4129 | -16.9649
9.4372 | -34.4133 | -16.975
9.4433 | -34.4138 | -16.985
9.4494 | -34.4143 | -16.995
9.4555 | -34.4147 | -17.0052
9.4616 | -34.4152 | -17.0152
9.4677 | -34.4157 | -17.0252
9.4738 | -34.4161 | -17.0353
9.4799 | -34.4166 | -17.0454
9.486 | -34.417 | -17.0555
9.4921 | -34.4175 | -17.0655
9.4982 | -34.418 | -17.0753
9.5043 | -34.4185 | -17.0854
9.5104 | -34.4189 | -17.0953
9.5165 | -34.4194 | -17.1053
9.5226 | -34.4198 | -17.1155
9.5288 | -34.4203 | -17.1255
9.5349 | -34.4208 | -17.1355
9.541 | -34.4212 | -17.1454
9.5471 | -34.4217 | -17.1554
9.5532 | -34.4221 | -17.1655
9.5593 | -34.4226 | -17.1757
9.5654 | -34.4231 | -17.1856
9.5715 | -34.4235 | -17.1956
9.5776 | -34.424 | -17.2058
9.5837 | -34.4244 | -17.2157
9.5898 | -34.4249 | -17.2259
9.5959 | -34.4253 | -17.236
9.602 | -34.4258 | -17.2458
9.6081 | -34.4263 | -17.2557
9.6142 | -34.4267 | -17.2657
9.6203 | -34.4272 | -17.2757
9.6264 | -34.4276 | -17.2857
9.6325 | -34.4281 | -17.2958
9.6386 | -34.4286 | -17.3057
9.6447 | -34.429 | -17.3158
9.6508 | -34.4295 | -17.3257
9.6569 | -34.43 | -17.3358
9.663 | -34.4304 | -17.3459
9.6691 | -34.4309 | -17.3558
9.6753 | -34.4314 | -17.3658
9.6814 | -34.4319 | -17.3757
9.6875 | -34.4323 | -17.3855
9.6936 | -34.4328 | -17.3953
9.6997 | -34.4333 | -17.4054
9.7058 | -34.4338 | -17.4154
9.7119 | -34.4343 | -17.4255
9.718 | -34.4347 | -17.4356
9.7241 | -34.4352 | -17.4455
9.7302 | -34.4357 | -17.4556
9.7363 | -34.4361 | -17.4655
9.7424 | -34.4366 | -17.4753
9.7485 | -34.4371 | -17.4853
9.7546 | -34.4375 | -17.4951
9.7607 | -34.438 | -17.5053
9.7668 | -34.4385 | -17.5154
9.7729 | -34.439 | -17.5254
9.779 | -34.4395 | -17.5351
9.7851 | -34.4399 | -17.5453
9.7912 | -34.4404 | -17.5551
9.7973 | -34.4409 | -17.5652
9.8034 | -34.4414 | -17.575
9.8095 | -34.4419 | -17.5852
9.8157 | -34.4424 | -17.595
9.8218 | -34.4428 | -17.6052
9.8279 | -34.4433 | -17.615
9.834 | -34.4438 | -17.6252
9.8401 | -34.4442 | -17.6352
9.8462 | -34.4447 | -17.6448
9.8523 | -34.4452 | -17.6548
9.8584 | -34.4457 | -17.6647
9.8645 | -34.4461 | -17.6749
9.8706 | -34.4466 | -17.6849
9.8767 | -34.447 | -17.6948
9.8828 | -34.4475 | -17.7046
9.8889 | -34.448 | -17.7144
9.895 | -34.4484 | -17.7244
9.9011 | -34.4489 | -17.7344
9.9072 | -34.4494 | -17.7443
9.9133 | -34.4498 | -17.7543
9.9194 | -34.4503 | -17.7641
9.9255 | -34.4508 | -17.774
9.9316 | -34.4512 | -17.784
9.9377 | -34.4517 | -17.7938
9.9438 | -34.4522 | -17.8038
9.9499 | -34.4527 | -17.8139
9.956 | -34.4531 | -17.8238
9.9622 | -34.4536 | -17.8339
9.9683 | -34.4541 | -17.8439
9.9744 | -34.4545 | -17.8539
9.9805 | -34.455 | -17.8639
9.9866 | -34.4555 | -17.8738
9.9927 | -34.4559 | -17.8836
9.9988 | -34.4564 | -17.8937
10.0049 | -34.4569 | -17.9036
10.011 | -34.4574 | -17.9136
10.0171 | -34.4578 | -17.9234
10.0232 | -34.4583 | -17.9331
10.0293 | -34.4588 | -17.9431
10.0354 | -34.4593 | -17.953
10.0415 | -34.4598 | -17.963
10.0476 | -34.4602 | -17.9731
10.0537 | -34.4607 | -17.9831
10.0598 | -34.4612 | -17.9929
10.0659 | -34.4616 | -18.0029
10.072 | -34.4621 | -18.0127
10.0781 | -34.4626 | -18.0225
10.0842 | -34.4631 | -18.0325
10.0903 | -34.4635 | -18.0424
10.0964 | -34.464 | -18.052
10.1026 | -34.4645 | -18.062
10.1087 | -34.4649 | -18.0719
10.1148 | -34.4653 | -18.0819
10.1209 | -34.4658 | -18.0917
10.127 | -34.4663 | -18.1014
10.1331 | -34.4668 | -18.1112
10.1392 | -34.4672 | -18.1209
10.1453 | -34.4677 | -18.1308
10.1514 | -34.4682 | -18.1407
10.1575 | -34.4687 | -18.1507
10.1636 | -34.4691 | -18.1604
10.1697 | -34.4696 | -18.1704
10.1758 | -34.4701 | -18.1801
10.1819 | -34.4706 | -18.1899
10.188 | -34.471 | -18.1998
10.1941 | -34.4715 | -18.2097
10.2002 | -34.4719 | -18.2195
10.2063 | -34.4724 | -18.2294
10.2124 | -34.4729 | -18.2393
10.2185 | -34.4733 | -18.2492
10.2246 | -34.4738 | -18.259
10.2307 | -34.4743 | -18.2689
10.2368 | -34.4748 | -18.2784
10.2429 | -34.4752 | -18.2883
10.2491 | -34.4757 | -18.2982
10.2552 | -34.4762 | -18.3082
10.2613 | -34.4766 | -18.3181
10.2674 | -34.4771 | -18.3279
10.2735 | -34.4776 | -18.3379
10.2796 | -34.4781 | -18.3478
10.2857 | -34.4786 | -18.3577
10.2918 | -34.479 | -18.3676
10.2979 | -34.4795 | -18.3775
10.304 | -34.48 | -18.3872
10.3101 | -34.4804 | -18.3971
10.3162 | -34.4809 | -18.4071
10.3223 | -34.4814 | -18.4168
10.3284 | -34.4819 | -18.4266
10.3345 | -34.4823 | -18.4363
10.3406 | -34.4828 | -18.4461
10.3467 | -34.4833 | -18.4559
10.3528 | -34.4837 | -18.4656
10.3589 | -34.4842 | -18.4755
10.365 | -34.4847 | -18.4855
10.3711 | -34.4851 | -18.4954
10.3772 | -34.4856 | -18.5052
10.3833 | -34.4861 | -18.515
10.3895 | -34.4866 | -18.5249
10.3956 | -34.4871 | -18.5347
10.4017 | -34.4876 | -18.5444
10.4078 | -34.488 | -18.5542
10.4139 | -34.4885 | -18.5639
10.42 | -34.489 | -18.5739
10.4261 | -34.4895 | -18.5838
10.4322 | -34.4899 | -18.5936
10.4383 | -34.4904 | -18.6034
10.4444 | -34.4908 | -18.6131
10.4505 | -34.4913 | -18.623
10.4566 | -34.4917 | -18.6328
10.4627 | -34.4922 | -18.6424
10.4688 | -34.4927 | -18.6522
10.4749 | -34.4931 | -18.6621
10.481 | -34.4936 | -18.6719
10.4871 | -34.4941 | -18.6817
10.4932 | -34.4946 | -18.6914
10.4993 | -34.4951 | -18.7012
10.5054 | -34.4955 | -18.711
10.5115 | -34.496 | -18.7207
10.5176 | -34.4965 | -18.7303
10.5237 | -34.497 | -18.7402
10.5298 | -34.4975 | -18.7498
10.536 | -34.4979 | -18.7597
10.5421 | -34.4984 | -18.7693
10.5482 | -34.4989 | -18.7792
10.5543 | -34.4993 | -18.7887
10.5604 | -34.4998 | -18.7983
10.5665 | -34.5003 | -18.8079
10.5726 | -34.5008 | -18.8177
10.5787 | -34.5012 | -18.8274
10.5848 | -34.5017 | -18.8368
10.5909 | -34.5022 | -18.8467
10.597 | -34.5026 | -18.8565
10.6031 | -34.5031 | -18.8664
10.6092 | -34.5036 | -18.8762
10.6153 | -34.504 | -18.8862
10.6214 | -34.5045 | -18.8956
10.6275 | -34.505 | -18.9053
10.6336 | -34.5054 | -18.9152
10.6397 | -34.5059 | -18.925
10.6458 | -34.5064 | -18.9346
10.6519 | -34.5068 | -18.9445
10.658 | -34.5073 | -18.9542
10.6641 | -34.5078 | -18.9641
10.6702 | -34.5083 | -18.9739
10.6764 | -34.5087 | -18.9838
10.6825 | -34.5092 | -18.9934
10.6886 | -34.5097 | -19.003
10.6947 | -34.5102 | -19.0126
10.7008 | -34.5107 | -19.0223
10.7069 | -34.5111 | -19.0321
10.713 | -34.5116 | -19.0419
10.7191 | -34.5121 | -19.0516
10.7252 | -34.5126 | -19.061
10.7313 | -34.513 | -19.0707
10.7374 | -34.5135 | -19.0805
10.7435 | -34.514 | -19.0903
10.7496 | -34.5144 | -19.0998
10.7557 | -34.5149 | -19.1097
10.7618 | -34.5154 | -19.1191
10.7679 | -34.5158 | -19.1288
10.774 | -34.5163 | -19.1384
10.7801 | -34.5168 | -19.1481
10.7862 | -34.5172 | -19.1576
10.7923 | -34.5177 | -19.1675
10.7984 | -34.5182 | -19.177
10.8045 | -34.5187 | -19.1865
10.8106 | -34.5191 | -19.1962
10.8168 | -34.5196 | -19.2057
10.8229 | -34.52 | -19.2154
10.829 | -34.5205 | -19.225
10.8351 | -34.521 | -19.2344
10.8412 | -34.5215 | -19.244
10.8473 | -34.5219 | -19.2534
10.8534 | -34.5224 | -19.2631
10.8595 | -34.5229 | -19.2728
10.8656 | -34.5233 | -19.2821
10.8717 | -34.5238 | -19.2918
10.8778 | -34.5243 | -19.3014
10.8839 | -34.5247 | -19.3111
10.89 | -34.5252 | -19.3208
10.8961 | -34.5257 | -19.3303
10.9022 | -34.5262 | -19.3399
10.9083 | -34.5266 | -19.3494
10.9144 | -34.5271 | -19.359
10.9205 | -34.5276 | -19.3685
10.9266 | -34.528 | -19.3782
10.9327 | -34.5285 | -19.3878
10.9388 | -34.529 | -19.3973
10.9449 | -34.5295 | -19.407
10.951 | -34.5299 | -19.4166
10.9571 | -34.5304 | -19.4261
10.9633 | -34.5309 | -19.4357
10.9694 | -34.5313 | -19.4455
10.9755 | -34.5318 | -19.4551
10.9816 | -34.5323 | -19.4648
10.9877 | -34.5327 | -19.4744
10.9938 | -34.5332 | -19.4841
10.9999 | -34.5337 | -19.4937
11.006 | -34.5342 | -19.5032
11.0121 | -34.5346 | -19.5128
11.0182 | -34.5351 | -19.5223
11.0243 | -34.5356 | -19.5319
11.0304 | -34.5361 | -19.5416
11.0365 | -34.5366 | -19.5513
11.0426 | -34.537 | -19.5606
11.0487 | -34.5376 | -19.5702
11.0548 | -34.5381 | -19.5798
11.0609 | -34.5385 | -19.5893
11.067 | -34.539 | -19.5989
11.0731 | -34.5395 | -19.6084
11.0792 | -34.54 | -19.618
11.0853 | -34.5405 | -19.6275
11.0914 | -34.541 | -19.6371
11.0975 | -34.5414 | -19.6466
11.1037 | -34.5419 | -19.6561
11.1098 | -34.5424 | -19.6659
11.1159 | -34.5429 | -19.6754
11.122 | -34.5433 | -19.6849
11.1281 | -34.5438 | -19.6946
11.1342 | -34.5443 | -19.7042
11.1403 | -34.5448 | -19.7139
11.1464 | -34.5453 | -19.7236
11.1525 | -34.5458 | -19.7328
11.1586 | -34.5462 | -19.7425
11.1647 | -34.5467 | -19.7518
11.1708 | -34.5471 | -19.7613
11.1769 | -34.5476 | -19.7709
11.183 | -34.5481 | -19.7805
11.1891 | -34.5486 | -19.7901
11.1952 | -34.5491 | -19.7996
11.2013 | -34.5496 | -19.8094
11.2074 | -34.55 | -19.8191
11.2135 | -34.5505 | -19.8289
11.2196 | -34.551 | -19.8382
11.2257 | -34.5515 | -19.8479
11.2318 | -34.552 | -19.8574
11.2379 | -34.5525 | -19.8671
11.244 | -34.553 | -19.8767
11.2502 | -34.5535 | -19.8862
11.2563 | -34.554 | -19.8957
11.2624 | -34.5544 | -19.905
11.2685 | -34.5549 | -19.9145
11.2746 | -34.5554 | -19.9238
11.2807 | -34.5559 | -19.9335
11.2868 | -34.5563 | -19.9431
11.2929 | -34.5568 | -19.952
11.299 | -34.5574 | -19.9618
11.3051 | -34.5578 | -19.9715
11.3112 | -34.5584 | -19.9813
11.3173 | -34.5589 | -19.9907
11.3234 | -34.5594 | -20.0002
11.3295 | -34.5599 | -20.0101
11.3356 | -34.5603 | -20.0196
11.3417 | -34.5608 | -20.0289
11.3478 | -34.5613 | -20.0385
11.3539 | -34.5617 | -20.0481
11.36 | -34.5623 | -20.0577
11.3661 | -34.5627 | -20.0669
11.3722 | -34.5633 | -20.0763
11.3783 | -34.5637 | -20.0859
11.3844 | -34.5642 | -20.0952
11.3906 | -34.5648 | -20.1046
11.3967 | -34.5652 | -20.1138
11.4028 | -34.5657 | -20.1234
11.4089 | -34.5662 | -20.1327
11.415 | -34.5667 | -20.1421
11.4211 | -34.5671 | -20.1518
11.4272 | -34.5676 | -20.1613
11.4333 | -34.5681 | -20.1708
11.4394 | -34.5686 | -20.1803
11.4455 | -34.569 | -20.1897
11.4516 | -34.5695 | -20.1992
11.4577 | -34.57 | -20.2088
11.4638 | -34.5705 | -20.218
11.4699 | -34.571 | -20.2276
11.476 | -34.5715 | -20.2373
11.4821 | -34.572 | -20.2468
11.4882 | -34.5725 | -20.2564
11.4943 | -34.573 | -20.266
11.5004 | -34.5735 | -20.2754
11.5065 | -34.574 | -20.2849
11.5126 | -34.5744 | -20.2942
11.5187 | -34.5749 | -20.3037
11.5248 | -34.5754 | -20.3133
11.5309 | -34.5759 | -20.3228
11.5371 | -34.5764 | -20.3321
11.5432 | -34.5769 | -20.3417
11.5493 | -34.5774 | -20.3513
11.5554 | -34.5778 | -20.3606
11.5615 | -34.5783 | -20.3702
11.5676 | -34.5789 | -20.3797
11.5737 | -34.5794 | -20.3891
11.5798 | -34.5799 | -20.3983
11.5859 | -34.5804 | -20.4079
11.592 | -34.5808 | -20.4174
11.5981 | -34.5813 | -20.4268
11.6042 | -34.5818 | -20.436
11.6103 | -34.5823 | -20.4453
11.6164 | -34.5828 | -20.4546
11.6225 | -34.5833 | -20.464
11.6286 | -34.5838 | -20.4735
11.6347 | -34.5843 | -20.4829
11.6408 | -34.5848 | -20.4923
11.6469 | -34.5853 | -20.5019
11.653 | -34.5858 | -20.5111
11.6591 | -34.5863 | -20.5205
11.6652 | -34.5867 | -20.5297
11.6713 | -34.5872 | -20.5391
11.6775 | -34.5877 | -20.5484
11.6836 | -34.5882 | -20.5576
11.6897 | -34.5887 | -20.5671
11.6958 | -34.5892 | -20.5766
11.7019 | -34.5897 | -20.5862
11.708 | -34.5902 | -20.5957
11.7141 | -34.5907 | -20.6049
11.7202 | -34.5912 | -20.6143
11.7263 | -34.5917 | -20.6238
11.7324 | -34.5922 | -20.6333
11.7385 | -34.5926 | -20.6427
11.7446 | -34.5931 | -20.6522
11.7507 | -34.5936 | -20.6613
11.7568 | -34.5941 | -20.6707
11.7629 | -34.5946 | -20.68
11.769 | -34.5951 | -20.6893
11.7751 | -34.5956 | -20.6987
11.7812 | -34.5961 | -20.708
11.7873 | -34.5966 | -20.7174
11.7934 | -34.5971 | -20.7268
11.7995 | -34.5976 | -20.7362
11.8056 | -34.5981 | -20.7455
11.8117 | -34.5986 | -20.7548
11.8178 | -34.5991 | -20.764
11.824 | -34.5996 | -20.7733
11.8301 | -34.6001 | -20.7827
11.8362 | -34.6006 | -20.7919
11.8423 | -34.6011 | -20.8013
11.8484 | -34.6016 | -20.8106
11.8545 | -34.6021 | -20.8198
11.8606 | -34.6026 | -20.8291
11.8667 | -34.6031 | -20.8386
11.8728 | -34.6036 | -20.8478
11.8789 | -34.6041 | -20.8571
11.885 | -34.6046 | -20.8665
11.8911 | -34.6051 | -20.8759
11.8972 | -34.6056 | -20.8852
11.9033 | -34.6061 | -20.8946
11.9094 | -34.6066 | -20.9038
11.9155 | -34.6071 | -20.9131
11.9216 | -34.6076 | -20.9223
11.9277 | -34.6081 | -20.9315
11.9338 | -34.6086 | -20.9409
11.9399 | -34.6091 | -20.9503
11.946 | -34.6096 | -20.9595
11.9521 | -34.6101 | -20.9688
11.9582 | -34.6106 | -20.9779
11.9644 | -34.6111 | -20.9872
11.9705 | -34.6116 | -20.9965
11.9766 | -34.6121 | -21.0058
11.9827 | -34.6126 | -21.0152
11.9888 | -34.6132 | -21.0245
11.9949 | -34.6137 | -21.0335
12.001 | -34.6142 | -21.0428
12.0071 | -34.6147 | -21.0521
12.0132 | -34.6152 | -21.0613
12.0193 | -34.6157 | -21.0707
12.0254 | -34.6161 | -21.08
12.0315 | -34.6167 | -21.0891
12.0376 | -34.6171 | -21.0983
12.0437 | -34.6176 | -21.1075
12.0498 | -34.6181 | -21.1168
12.0559 | -34.6186 | -21.126
12.062 | -34.6192 | -21.1352
12.0681 | -34.6197 | -21.1446
12.0742 | -34.6202 | -21.1537
12.0803 | -34.6207 | -21.1632
12.0864 | -34.6212 | -21.1725
12.0925 | -34.6217 | -21.1817
12.0986 | -34.6222 | -21.1908
12.1047 | -34.6227 | -21.2001
12.1109 | -34.6233 | -21.2092
12.117 | -34.6238 | -21.2184
12.1231 | -34.6243 | -21.2276
12.1292 | -34.6248 | -21.2368
12.1353 | -34.6253 | -21.246
12.1414 | -34.6258 | -21.2553
12.1475 | -34.6263 | -21.2647
12.1536 | -34.6268 | -21.274
12.1597 | -34.6273 | -21.2833
12.1658 | -34.6278 | -21.2924
12.1719 | -34.6283 | -21.3014
12.178 | -34.6288 | -21.3107
12.1841 | -34.6294 | -21.3199
12.1902 | -34.6299 | -21.3292
12.1963 | -34.6304 | -21.3385
12.2024 | -34.6309 | -21.3476
12.2085 | -34.6314 | -21.3569
12.2146 | -34.6319 | -21.3661
12.2207 | -34.6324 | -21.3753
12.2268 | -34.6329 | -21.3845
12.2329 | -34.6334 | -21.3938
12.239 | -34.634 | -21.4027
12.2451 | -34.6345 | -21.412
12.2513 | -34.635 | -21.4213
12.2574 | -34.6355 | -21.4306
12.2635 | -34.636 | -21.4397
12.2696 | -34.6365 | -21.4488
12.2757 | -34.6371 | -21.4578
12.2818 | -34.6376 | -21.467
12.2879 | -34.6381 | -21.4762
12.294 | -34.6386 | -21.4853
12.3001 | -34.6391 | -21.4944
12.3062 | -34.6397 | -21.5036
12.3123 | -34.6402 | -21.5124
12.3184 | -34.6407 | -21.5215
12.3245 | -34.6412 | -21.5306
12.3306 | -34.6417 | -21.5397
12.3367 | -34.6422 | -21.5489
12.3428 | -34.6427 | -21.5582
12.3489 | -34.6433 | -21.567
12.355 | -34.6438 | -21.5763
12.3611 | -34.6443 | -21.5855
12.3672 | -34.6448 | -21.5946
12.3733 | -34.6453 | -21.6038
12.3794 | -34.6459 | -21.6126
12.3855 | -34.6464 | -21.6217
12.3916 | -34.6469 | -21.6308
12.3978 | -34.6474 | -21.6401
12.4039 | -34.6479 | -21.6492
12.41 | -34.6485 | -21.6583
12.4161 | -34.649 | -21.6672
12.4222 | -34.6495 | -21.6762
12.4283 | -34.65 | -21.6852
12.4344 | -34.6506 | -21.6946
12.4405 | -34.6511 | -21.7038
12.4466 | -34.6516 | -21.7129
12.4527 | -34.6521 | -21.7219
12.4588 | -34.6526 | -21.731
12.4649 | -34.6531 | -21.7401
12.471 | -34.6536 | -21.7493
12.4771 | -34.6541 | -21.7583
12.4832 | -34.6546 | -21.7673
12.4893 | -34.6552 | -21.7763
12.4954 | -34.6557 | -21.7853
12.5015 | -34.6562 | -21.7944
12.5076 | -34.6567 | -21.8035
12.5137 | -34.6572 | -21.8126
12.5198 | -34.6578 | -21.8215
12.5259 | -34.6583 | -21.8307
12.532 | -34.6588 | -21.8398
12.5382 | -34.6593 | -21.849
12.5443 | -34.6598 | -21.858
12.5504 | -34.6603 | -21.8672
12.5565 | -34.6608 | -21.8761
12.5626 | -34.6614 | -21.8852
12.5687 | -34.6619 | -21.8943
12.5748 | -34.6624 | -21.9033
12.5809 | -34.6629 | -21.9124
12.587 | -34.6634 | -21.9216
12.5931 | -34.6639 | -21.9304
12.5992 | -34.6645 | -21.9396
12.6053 | -34.665 | -21.9486
12.6114 | -34.6655 | -21.9575
12.6175 | -34.666 | -21.9667
12.6236 | -34.6665 | -21.9758
12.6297 | -34.6671 | -21.9848
12.6358 | -34.6676 | -21.9937
12.6419 | -34.6681 | -22.0028
12.648 | -34.6686 | -22.0119
12.6541 | -34.6691 | -22.0209
12.6602 | -34.6697 | -22.0296
12.6663 | -34.6702 | -22.0388
12.6724 | -34.6707 | -22.0477
12.6785 | -34.6712 | -22.0568
12.6847 | -34.6718 | -22.0658
12.6908 | -34.6723 | -22.0748
12.6969 | -34.6728 | -22.0838
12.703 | -34.6733 | -22.093
12.7091 | -34.6738 | -22.102
12.7152 | -34.6743 | -22.111
12.7213 | -34.6748 | -22.1201
12.7274 | -34.6754 | -22.1291
12.7335 | -34.6759 | -22.1379
12.7396 | -34.6764 | -22.1467
12.7457 | -34.677 | -22.1558
12.7518 | -34.6775 | -22.1649
12.7579 | -34.678 | -22.174
12.764 | -34.6785 | -22.1831
12.7701 | -34.679 | -22.192
12.7762 | -34.6796 | -22.2009
12.7823 | -34.6801 | -22.21
12.7884 | -34.6806 | -22.2191
12.7945 | -34.6811 | -22.2283
12.8006 | -34.6817 | -22.2373
12.8067 | -34.6822 | -22.2462
12.8128 | -34.6827 | -22.2551
12.8189 | -34.6832 | -22.264
12.8251 | -34.6837 | -22.2731
12.8312 | -34.6843 | -22.2822
12.8373 | -34.6848 | -22.2912
12.8434 | -34.6853 | -22.3002
12.8495 | -34.6858 | -22.3092
12.8556 | -34.6864 | -22.3182
12.8617 | -34.6869 | -22.3273
12.8678 | -34.6874 | -22.3364
12.8739 | -34.6879 | -22.3452
12.88 | -34.6884 | -22.3541
12.8861 | -34.6889 | -22.363
12.8922 | -34.6895 | -22.372
12.8983 | -34.69 | -22.381
12.9044 | -34.6905 | -22.3899
12.9105 | -34.691 | -22.3987
12.9166 | -34.6916 | -22.4077
12.9227 | -34.6921 | -22.4167
12.9288 | -34.6926 | -22.4257
12.9349 | -34.6931 | -22.4346
12.941 | -34.6937 | -22.4435
12.9471 | -34.6942 | -22.4526
12.9532 | -34.6948 | -22.4615
12.9593 | -34.6953 | -22.4707
12.9654 | -34.6958 | -22.4797
12.9716 | -34.6963 | -22.4887
12.9777 | -34.6969 | -22.4975
12.9838 | -34.6974 | -22.5065
12.9899 | -34.6979 | -22.5154
12.996 | -34.6984 | -22.5243
13.0021 | -34.6989 | -22.5332
13.0082 | -34.6995 | -22.5421
13.0143 | -34.7 | -22.5509
13.0204 | -34.7005 | -22.5597
13.0265 | -34.7011 | -22.5686
13.0326 | -34.7016 | -22.5776
13.0387 | -34.7021 | -22.5868
13.0448 | -34.7026 | -22.5958
13.0509 | -34.7031 | -22.6046
13.057 | -34.7037 | -22.6134
13.0631 | -34.7042 | -22.6223
13.0692 | -34.7047 | -22.6312
13.0753 | -34.7053 | -22.6403
13.0814 | -34.7058 | -22.6494
13.0875 | -34.7063 | -22.6582
13.0936 | -34.7068 | -22.6671
13.0997 | -34.7074 | -22.6759
13.1058 | -34.7079 | -22.6848
13.112 | -34.7084 | -22.6937
13.1181 | -34.709 | -22.7025
13.1242 | -34.7095 | -22.7114
13.1303 | -34.71 | -22.7203
13.1364 | -34.7105 | -22.7292
13.1425 | -34.711 | -22.7381
13.1486 | -34.7115 | -22.747
13.1547 | -34.7121 | -22.7558
13.1608 | -34.7126 | -22.7647
13.1669 | -34.7131 | -22.7734
13.173 | -34.7136 | -22.7824
13.1791 | -34.7141 | -22.7912
13.1852 | -34.7147 | -22.8001
13.1913 | -34.7152 | -22.8089
13.1974 | -34.7157 | -22.8179
13.2035 | -34.7162 | -22.8268
13.2096 | -34.7167 | -22.8355
13.2157 | -34.7173 | -22.8442
13.2218 | -34.7178 | -22.8532
13.2279 | -34.7183 | -22.862
13.234 | -34.7189 | -22.8708
13.2401 | -34.7194 | -22.8797
13.2462 | -34.7199 | -22.8885
13.2524 | -34.7204 | -22.8974
13.2585 | -34.7209 | -22.9062
13.2646 | -34.7215 | -22.9151
13.2707 | -34.722 | -22.924
13.2768 | -34.7225 | -22.933
13.2829 | -34.7231 | -22.942
13.289 | -34.7236 | -22.9507
13.2951 | -34.7241 | -22.9595
13.3012 | -34.7247 | -22.9684
13.3073 | -34.7252 | -22.9773
13.3134 | -34.7257 | -22.9863
13.3195 | -34.7262 | -22.9951
13.3256 | -34.7268 | -23.0039
13.3317 | -34.7273 | -23.0126
13.3378 | -34.7278 | -23.0214
13.3439 | -34.7284 | -23.0304
13.35 | -34.7289 | -23.0393
13.3561 | -34.7294 | -23.0481
13.3622 | -34.7299 | -23.0569
13.3683 | -34.7305 | -23.0657
13.3744 | -34.731 | -23.0745
13.3805 | -34.7315 | -23.0834
13.3866 | -34.732 | -23.092
13.3927 | -34.7325 | -23.1009
13.3989 | -34.733 | -23.1097
13.405 | -34.7336 | -23.1184
13.4111 | -34.7341 | -23.1272
13.4172 | -34.7346 | -23.1361
13.4233 | -34.7352 | -23.1451
13.4294 | -34.7357 | -23.1542
13.4355 | -34.7362 | -23.1628
13.4416 | -34.7367 | -23.1717
13.4477 | -34.7372 | -23.1806
13.4538 | -34.7378 | -23.1895
13.4599 | -34.7383 | -23.1984
13.466 | -34.7388 | -23.2072
13.4721 | -34.7393 | -23.216
13.4782 | -34.7398 | -23.2248
13.4843 | -34.7404 | -23.2338
13.4904 | -34.7409 | -23.2425
13.4965 | -34.7414 | -23.2516
13.5026 | -34.7419 | -23.2604
13.5087 | -34.7425 | -23.2693
13.5148 | -34.743 | -23.2783
13.5209 | -34.7435 | -23.2872
13.527 | -34.744 | -23.296
13.5331 | -34.7445 | -23.3047
13.5393 | -34.7451 | -23.3134
13.5454 | -34.7456 | -23.3222
13.5515 | -34.7461 | -23.3312
13.5576 | -34.7466 | -23.34
13.5637 | -34.7472 | -23.3489
13.5698 | -34.7477 | -23.3578
13.5759 | -34.7482 | -23.3663
13.582 | -34.7487 | -23.3749
13.5881 | -34.7493 | -23.3838
13.5942 | -34.7498 | -23.3927
13.6003 | -34.7503 | -23.4015
13.6064 | -34.7508 | -23.4104
13.6125 | -34.7514 | -23.4193
13.6186 | -34.7519 | -23.4281
13.6247 | -34.7524 | -23.437
13.6308 | -34.7529 | -23.4458
13.6369 | -34.7534 | -23.4547
13.643 | -34.754 | -23.4636
13.6491 | -34.7545 | -23.4722
13.6552 | -34.755 | -23.4811
13.6613 | -34.7555 | -23.49
13.6674 | -34.756 | -23.4989
13.6735 | -34.7566 | -23.5077
13.6796 | -34.7571 | -23.5165
13.6858 | -34.7576 | -23.5252
13.6919 | -34.7581 | -23.534
13.698 | -34.7586 | -23.5429
13.7041 | -34.7592 | -23.5518
13.7102 | -34.7597 | -23.5606
13.7163 | -34.7602 | -23.5694
13.7224 | -34.7608 | -23.5782
13.7285 | -34.7613 | -23.587
13.7346 | -34.7618 | -23.5959
13.7407 | -34.7624 | -23.6047
13.7468 | -34.7629 | -23.6135
13.7529 | -34.7634 | -23.6221
13.759 | -34.7639 | -23.631
13.7651 | -34.7645 | -23.6398
13.7712 | -34.765 | -23.6486
13.7773 | -34.7655 | -23.6575
13.7834 | -34.766 | -23.6663
13.7895 | -34.7666 | -23.675
13.7956 | -34.7671 | -23.684
13.8017 | -34.7676 | -23.6927
13.8078 | -34.7681 | -23.7015
13.8139 | -34.7686 | -23.7102
13.82 | -34.7691 | -23.7189
13.8262 | -34.7697 | -23.7278
13.8323 | -34.7702 | -23.7366
13.8384 | -34.7707 | -23.7454
13.8445 | -34.7712 | -23.7542
13.8506 | -34.7718 | -23.7631
13.8567 | -34.7723 | -23.7721
13.8628 | -34.7728 | -23.781
13.8689 | -34.7734 | -23.7899
13.875 | -34.7739 | -23.7986
13.8811 | -34.7744 | -23.8074
13.8872 | -34.7749 | -23.8161
13.8933 | -34.7754 | -23.8248
13.8994 | -34.7759 | -23.8336
13.9055 | -34.7765 | -23.8425
13.9116 | -34.777 | -23.8513
13.9177 | -34.7775 | -23.8602
13.9238 | -34.778 | -23.8689
13.9299 | -34.7785 | -23.8776
13.936 | -34.779 | -23.8865
13.9421 | -34.7796 | -23.8952
13.9482 | -34.7801 | -23.904
13.9543 | -34.7806 | -23.9129
13.9604 | -34.7811 | -23.9216
13.9665 | -34.7817 | -23.9303
13.9727 | -34.7822 | -23.9392
13.9788 | -34.7827 | -23.9479
13.9849 | -34.7832 | -23.9567
13.991 | -34.7838 | -23.9655
13.9971 | -34.7843 | -23.9741
14.0032 | -34.7848 | -23.9829
14.0093 | -34.7854 | -23.9916
14.0154 | -34.7859 | -24.0003
14.0215 | -34.7864 | -24.009
14.0276 | -34.7869 | -24.0178
14.0337 | -34.7875 | -24.0265
14.0398 | -34.788 | -24.0352
14.0459 | -34.7885 | -24.044
14.052 | -34.789 | -24.0528
14.0581 | -34.7895 | -24.0616
14.0642 | -34.7901 | -24.0704
14.0703 | -34.7906 | -24.0792
14.0764 | -34.7911 | -24.0878
14.0825 | -34.7917 | -24.0965
14.0886 | -34.7922 | -24.1052
14.0947 | -34.7927 | -24.114
14.1008 | -34.7933 | -24.1229
14.1069 | -34.7938 | -24.1315
14.1131 | -34.7943 | -24.1404
14.1192 | -34.7949 | -24.149
14.1253 | -34.7954 | -24.1578
14.1314 | -34.7959 | -24.1665
14.1375 | -34.7964 | -24.175
14.1436 | -34.797 | -24.1838
14.1497 | -34.7975 | -24.1924
14.1558 | -34.798 | -24.2012
14.1619 | -34.7985 | -24.2099
14.168 | -34.7991 | -24.2185
14.1741 | -34.7996 | -24.2271
14.1802 | -34.8001 | -24.2358
14.1863 | -34.8007 | -24.2446
14.1924 | -34.8012 | -24.2534
14.1985 | -34.8017 | -24.262
14.2046 | -34.8023 | -24.2709
14.2107 | -34.8028 | -24.2795
14.2168 | -34.8033 | -24.2882
14.2229 | -34.8038 | -24.2969
14.229 | -34.8044 | -24.3057
14.2351 | -34.8049 | -24.3144
14.2412 | -34.8054 | -24.3232
14.2473 | -34.806 | -24.3317
14.2534 | -34.8065 | -24.3403
14.2596 | -34.807 | -24.3491
14.2657 | -34.8075 | -24.3578
14.2718 | -34.8081 | -24.3665
14.2779 | -34.8086 | -24.3752
14.284 | -34.8092 | -24.384
14.2901 | -34.8097 | -24.3927
14.2962 | -34.8102 | -24.4014
14.3023 | -34.8107 | -24.4103
14.3084 | -34.8113 | -24.419
14.3145 | -34.8118 | -24.4278
14.3206 | -34.8123 | -24.4366
14.3267 | -34.8129 | -24.4453
14.3328 | -34.8134 | -24.454
14.3389 | -34.8139 | -24.4626
14.345 | -34.8144 | -24.4714
14.3511 | -34.815 | -24.4801
14.3572 | -34.8155 | -24.4889
14.3633 | -34.8161 | -24.4978
14.3694 | -34.8166 | -24.5064
14.3755 | -34.8172 | -24.5151
14.3816 | -34.8177 | -24.524
14.3877 | -34.8183 | -24.5326
14.3938 | -34.8188 | -24.5415
14.4 | -34.8193 | -24.5501
14.4061 | -34.8198 | -24.5589
14.4122 | -34.8204 | -24.5676
14.4183 | -34.8209 | -24.5762
14.4244 | -34.8214 | -24.585
14.4305 | -34.822 | -24.5937
14.4366 | -34.8225 | -24.6024
14.4427 | -34.8231 | -24.6112
14.4488 | -34.8236 | -24.6199
14.4549 | -34.8242 | -24.6285
14.461 | -34.8247 | -24.637
14.4671 | -34.8252 | -24.6457
14.4732 | -34.8258 | -24.6545
14.4793 | -34.8263 | -24.6633
14.4854 | -34.8269 | -24.6719
14.4915 | -34.8274 | -24.6804
14.4976 | -34.8279 | -24.6891
14.5037 | -34.8285 | -24.6977
14.5098 | -34.829 | -24.7064
14.5159 | -34.8295 | -24.7151
14.522 | -34.8301 | -24.7238
14.5281 | -34.8306 | -24.7323
14.5342 | -34.8312 | -24.7411
14.5403 | -34.8317 | -24.7497
14.5465 | -34.8322 | -24.7584
14.5526 | -34.8328 | -24.7672
14.5587 | -34.8333 | -24.7759
14.5648 | -34.8339 | -24.7844
14.5709 | -34.8344 | -24.7931
14.577 | -34.835 | -24.8018
14.5831 | -34.8355 | -24.8105
14.5892 | -34.836 | -24.8193
14.5953 | -34.8366 | -24.8278
14.6014 | -34.8371 | -24.8364
14.6075 | -34.8376 | -24.8451
14.6136 | -34.8381 | -24.8538
14.6197 | -34.8387 | -24.8625
14.6258 | -34.8392 | -24.8712
14.6319 | -34.8397 | -24.8798
14.638 | -34.8403 | -24.8885
14.6441 | -34.8408 | -24.8972
14.6502 | -34.8413 | -24.9059
14.6563 | -34.8419 | -24.9146
14.6624 | -34.8425 | -24.9232
14.6685 | -34.843 | -24.9316
14.6746 | -34.8436 | -24.9404
14.6807 | -34.8441 | -24.9489
14.6869 | -34.8447 | -24.9576
14.693 | -34.8452 | -24.9663
14.6991 | -34.8457 | -24.9749
14.7052 | -34.8463 | -24.9833
14.7113 | -34.8469 | -24.9921
14.7174 | -34.8474 | -25.0007
14.7235 | -34.8479 | -25.0092
14.7296 | -34.8485 | -25.0178
14.7357 | -34.8491 | -25.0263
14.7418 | -34.8496 | -25.035
14.7479 | -34.8501 | -25.0436
14.754 | -34.8507 | -25.0523
14.7601 | -34.8512 | -25.0608
14.7662 | -34.8518 | -25.0694
14.7723 | -34.8523 | -25.0778
14.7784 | -34.8529 | -25.0865
14.7845 | -34.8534 | -25.0951
14.7906 | -34.854 | -25.1039
14.7967 | -34.8545 | -25.1125
14.8028 | -34.855 | -25.1212
14.8089 | -34.8556 | -25.1297
14.815 | -34.8562 | -25.1382
14.8211 | -34.8567 | -25.1468
14.8272 | -34.8572 | -25.1556
14.8334 | -34.8578 | -25.1642
14.8395 | -34.8583 | -25.1726
14.8456 | -34.8589 | -25.1812
14.8517 | -34.8595 | -25.1899
14.8578 | -34.86 | -25.1986
14.8639 | -34.8606 | -25.2072
14.87 | -34.8611 | -25.2159
14.8761 | -34.8617 | -25.2244
14.8822 | -34.8622 | -25.2331
14.8883 | -34.8628 | -25.2418
14.8944 | -34.8633 | -25.2503
14.9005 | -34.8639 | -25.2589
14.9066 | -34.8644 | -25.2675
14.9127 | -34.865 | -25.2759
14.9188 | -34.8656 | -25.2845
14.9249 | -34.8661 | -25.2931
14.931 | -34.8666 | -25.3018
14.9371 | -34.8672 | -25.3103
14.9432 | -34.8678 | -25.319
14.9493 | -34.8683 | -25.3274
14.9554 | -34.8689 | -25.336
14.9615 | -34.8694 | -25.3446
14.9676 | -34.8699 | -25.3534
14.9738 | -34.8705 | -25.3619
14.9799 | -34.8711 | -25.3707
14.986 | -34.8716 | -25.3792
14.9921 | -34.8721 | -25.3878
14.9982 | -34.8727 | -25.3964
15.0043 | -34.8732 | -25.405
15.0104 | -34.8738 | -25.4136
15.0165 | -34.8744 | -25.4223
15.0226 | -34.8749 | -25.4307
15.0287 | -34.8755 | -25.4393
15.0348 | -34.8761 | -25.4479
15.0409 | -34.8766 | -25.4566
15.047 | -34.8772 | -25.4653
15.0531 | -34.8778 | -25.4737
15.0592 | -34.8783 | -25.4823
15.0653 | -34.8789 | -25.491
15.0714 | -34.8794 | -25.4995
15.0775 | -34.88 | -25.5079
15.0836 | -34.8805 | -25.5164
15.0897 | -34.8811 | -25.5246
15.0958 | -34.8817 | -25.5331
15.1019 | -34.8823 | -25.5417
15.108 | -34.8828 | -25.5502
15.1141 | -34.8834 | -25.5587
15.1203 | -34.8839 | -25.5673
15.1264 | -34.8845 | -25.5758
15.1325 | -34.8851 | -25.5844
15.1386 | -34.8856 | -25.5929
15.1447 | -34.8862 | -25.6014
15.1508 | -34.8867 | -25.61
15.1569 | -34.8873 | -25.6184
15.163 | -34.8879 | -25.6267
15.1691 | -34.8885 | -25.6352
15.1752 | -34.889 | -25.6436
15.1813 | -34.8896 | -25.6521
15.1874 | -34.8901 | -25.6607
15.1935 | -34.8907 | -25.669
15.1996 | -34.8912 | -25.6775
15.2057 | -34.8918 | -25.686
15.2118 | -34.8924 | -25.6945
15.2179 | -34.893 | -25.7031
15.224 | -34.8935 | -25.7115
15.2301 | -34.8941 | -25.7198
15.2362 | -34.8946 | -25.7284
15.2423 | -34.8952 | -25.737
15.2484 | -34.8958 | -25.7454
15.2545 | -34.8963 | -25.7539
15.2607 | -34.8969 | -25.7623
15.2668 | -34.8975 | -25.7706
15.2729 | -34.898 | -25.779
15.279 | -34.8986 | -25.7877
15.2851 | -34.8992 | -25.7961
15.2912 | -34.8997 | -25.8047
15.2973 | -34.9003 | -25.8132
15.3034 | -34.9009 | -25.8216
15.3095 | -34.9015 | -25.8301
15.3156 | -34.9021 | -25.8386
15.3217 | -34.9026 | -25.847
15.3278 | -34.9032 | -25.8557
15.3339 | -34.9038 | -25.8639
15.34 | -34.9044 | -25.8722
15.3461 | -34.9049 | -25.8807
15.3522 | -34.9055 | -25.8893
15.3583 | -34.9061 | -25.8979
15.3644 | -34.9066 | -25.9062
15.3705 | -34.9072 | -25.9144
15.3766 | -34.9078 | -25.9231
15.3827 | -34.9084 | -25.9314
15.3888 | -34.9089 | -25.94
15.3949 | -34.9095 | -25.9486
15.401 | -34.9101 | -25.957
15.4072 | -34.9107 | -25.9652
15.4133 | -34.9112 | -25.9737
15.4194 | -34.9118 | -25.9822
15.4255 | -34.9123 | -25.9907
15.4316 | -34.9129 | -25.9991
15.4377 | -34.9135 | -26.0078
15.4438 | -34.9141 | -26.0163
15.4499 | -34.9147 | -26.0247
15.456 | -34.9152 | -26.033
15.4621 | -34.9158 | -26.0414
15.4682 | -34.9164 | -26.0498
15.4743 | -34.917 | -26.058
15.4804 | -34.9175 | -26.0665
15.4865 | -34.9181 | -26.0751
15.4926 | -34.9187 | -26.0835
15.4987 | -34.9192 | -26.0918
15.5048 | -34.9198 | -26.1003
15.5109 | -34.9204 | -26.1084
15.517 | -34.921 | -26.1169
15.5231 | -34.9216 | -26.1254
15.5292 | -34.9221 | -26.1336
15.5353 | -34.9227 | -26.1421
15.5414 | -34.9232 | -26.1505
15.5476 | -34.9239 | -26.1585
15.5537 | -34.9244 | -26.1669
15.5598 | -34.925 | -26.1753
15.5659 | -34.9256 | -26.1837
15.572 | -34.9262 | -26.1922
15.5781 | -34.9267 | -26.2006
15.5842 | -34.9273 | -26.2088
15.5903 | -34.9279 | -26.2171
15.5964 | -34.9285 | -26.2253
15.6025 | -34.929 | -26.2337
15.6086 | -34.9296 | -26.2421
15.6147 | -34.9302 | -26.2506
15.6208 | -34.9308 | -26.2586
15.6269 | -34.9313 | -26.267
15.633 | -34.9319 | -26.2753
15.6391 | -34.9325 | -26.2837
15.6452 | -34.9331 | -26.2921
15.6513 | -34.9337 | -26.3002
15.6574 | -34.9343 | -26.3086
15.6635 | -34.9349 | -26.317
15.6696 | -34.9354 | -26.3255
15.6757 | -34.936 | -26.3338
15.6818 | -34.9365 | -26.342
15.688 | -34.9371 | -26.3501
15.6941 | -34.9377 | -26.3584
15.7002 | -34.9383 | -26.3668
15.7063 | -34.9388 | -26.3752
15.7124 | -34.9393 | -26.3835
15.7185 | -34.9399 | -26.392
15.7246 | -34.9405 | -26.4002
15.7307 | -34.9411 | -26.4085
15.7368 | -34.9417 | -26.4168
15.7429 | -34.9422 | -26.4251
15.749 | -34.9428 | -26.4334
15.7551 | -34.9434 | -26.4419
15.7612 | -34.944 | -26.45
15.7673 | -34.9446 | -26.4584
15.7734 | -34.9452 | -26.4668
15.7795 | -34.9457 | -26.4751
15.7856 | -34.9463 | -26.4833
15.7917 | -34.9469 | -26.4914
15.7978 | -34.9475 | -26.4998
15.8039 | -34.9481 | -26.5081
15.81 | -34.9486 | -26.5164
15.8161 | -34.9492 | -26.5245
15.8222 | -34.9498 | -26.5327
15.8283 | -34.9504 | -26.5408
15.8345 | -34.9509 | -26.5491
15.8406 | -34.9515 | -26.5573
15.8467 | -34.9521 | -26.5658
15.8528 | -34.9527 | -26.5741
15.8589 | -34.9533 | -26.5823
15.865 | -34.9539 | -26.5903
15.8711 | -34.9545 | -26.5988
15.8772 | -34.9551 | -26.6071
15.8833 | -34.9557 | -26.6155
15.8894 | -34.9562 | -26.6238
15.8955 | -34.9568 | -26.6319
15.9016 | -34.9574 | -26.64
15.9077 | -34.9579 | -26.6482
15.9138 | -34.9585 | -26.6565
15.9199 | -34.9591 | -26.6646
15.926 | -34.9597 | -26.6727
15.9321 | -34.9603 | -26.6808
15.9382 | -34.9609 | -26.689
15.9443 | -34.9614 | -26.6973
15.9504 | -34.962 | -26.7056
15.9565 | -34.9626 | -26.7138
15.9626 | -34.9632 | -26.7222
15.9687 | -34.9637 | -26.7302
15.9749 | -34.9643 | -26.7384
15.981 | -34.9649 | -26.7467
15.9871 | -34.9655 | -26.7551
15.9932 | -34.9661 | -26.7632
15.9993 | -34.9666 | -26.7715
16.0054 | -34.9672 | -26.7793
16.0115 | -34.9678 | -26.7877
16.0176 | -34.9684 | -26.796
16.0237 | -34.969 | -26.8044
16.0298 | -34.9696 | -26.8125
16.0359 | -34.9701 | -26.8209
16.042 | -34.9708 | -26.8289
16.0481 | -34.9713 | -26.8373
16.0542 | -34.9719 | -26.8457
16.0603 | -34.9725 | -26.8536
16.0664 | -34.973 | -26.8619
16.0725 | -34.9736 | -26.8698
16.0786 | -34.9742 | -26.8779
16.0847 | -34.9747 | -26.886
16.0908 | -34.9753 | -26.8942
16.0969 | -34.9759 | -26.9024
16.103 | -34.9765 | -26.9105
16.1091 | -34.9771 | -26.9185
16.1152 | -34.9776 | -26.9266
16.1214 | -34.9782 | -26.9349
16.1275 | -34.9788 | -26.9432
16.1336 | -34.9794 | -26.9512
16.1397 | -34.98 | -26.9593
16.1458 | -34.9806 | -26.9673
16.1519 | -34.9811 | -26.9756
16.158 | -34.9817 | -26.9839
16.1641 | -34.9823 | -26.9922
16.1702 | -34.9829 | -27.0004
16.1763 | -34.9834 | -27.0088
16.1824 | -34.984 | -27.0168
16.1885 | -34.9846 | -27.025
16.1946 | -34.9852 | -27.0331
16.2007 | -34.9858 | -27.0415
16.2068 | -34.9864 | -27.0497
16.2129 | -34.9869 | -27.058
16.219 | -34.9875 | -27.0658
16.2251 | -34.9881 | -27.0739
16.2312 | -34.9887 | -27.082
16.2373 | -34.9893 | -27.09
16.2434 | -34.9899 | -27.0982
16.2495 | -34.9905 | -27.106
16.2556 | -34.9911 | -27.1142
16.2618 | -34.9917 | -27.1222
16.2679 | -34.9922 | -27.1304
16.274 | -34.9928 | -27.1386
16.2801 | -34.9934 | -27.1466
16.2862 | -34.994 | -27.1546
16.2923 | -34.9946 | -27.1629
16.2984 | -34.9951 | -27.171
16.3045 | -34.9957 | -27.179
16.3106 | -34.9963 | -27.1872
16.3167 | -34.9968 | -27.1953
16.3228 | -34.9974 | -27.2033
16.3289 | -34.998 | -27.2114
16.335 | -34.9986 | -27.2194
16.3411 | -34.9992 | -27.2273
16.3472 | -34.9998 | -27.2354
16.3533 | -35.0003 | -27.2435
16.3594 | -35.0009 | -27.2513
16.3655 | -35.0015 | -27.2595
16.3716 | -35.0021 | -27.2676
16.3777 | -35.0027 | -27.2756
16.3838 | -35.0032 | -27.2836
16.3899 | -35.0038 | -27.2915
16.396 | -35.0044 | -27.2997
16.4021 | -35.005 | -27.3077
16.4083 | -35.0056 | -27.3158
16.4144 | -35.0062 | -27.3239
16.4205 | -35.0068 | -27.3322
16.4266 | -35.0073 | -27.3401
16.4327 | -35.0079 | -27.3481
16.4388 | -35.0085 | -27.3562
16.4449 | -35.0091 | -27.3643
16.451 | -35.0096 | -27.3724
16.4571 | -35.0102 | -27.3807
16.4632 | -35.0108 | -27.3886
16.4693 | -35.0113 | -27.3966
16.4754 | -35.0119 | -27.4047
16.4815 | -35.0125 | -27.4128
16.4876 | -35.0131 | -27.4209
16.4937 | -35.0137 | -27.429
16.4998 | -35.0143 | -27.4368
16.5059 | -35.0149 | -27.4449
16.512 | -35.0154 | -27.453
16.5181 | -35.016 | -27.4613
16.5242 | -35.0166 | -27.4695
16.5303 | -35.0172 | -27.4775
16.5364 | -35.0177 | -27.4855
16.5425 | -35.0183 | -27.4935
16.5487 | -35.0189 | -27.5014
16.5548 | -35.0195 | -27.5095
16.5609 | -35.0201 | -27.5176
16.567 | -35.0207 | -27.5254
16.5731 | -35.0212 | -27.5335
16.5792 | -35.0218 | -27.5416
16.5853 | -35.0224 | -27.5497
16.5914 | -35.023 | -27.5579
16.5975 | -35.0236 | -27.566
16.6036 | -35.0242 | -27.5738
16.6097 | -35.0247 | -27.5821
16.6158 | -35.0253 | -27.59
16.6219 | -35.0259 | -27.5981
16.628 | -35.0265 | -27.6061
16.6341 | -35.0271 | -27.6142
16.6402 | -35.0276 | -27.6219
16.6463 | -35.0282 | -27.63
16.6524 | -35.0288 | -27.6378
16.6585 | -35.0294 | -27.646
16.6646 | -35.03 | -27.6538
16.6707 | -35.0306 | -27.6617
16.6768 | -35.0312 | -27.6697
16.6829 | -35.0317 | -27.6777
16.689 | -35.0323 | -27.6857
16.6952 | -35.0329 | -27.6938
16.7013 | -35.0334 | -27.7018
16.7074 | -35.034 | -27.7095
16.7135 | -35.0346 | -27.7175
16.7196 | -35.0352 | -27.7255
16.7257 | -35.0358 | -27.7335
16.7318 | -35.0363 | -27.7415
16.7379 | -35.0369 | -27.7494
16.744 | -35.0375 | -27.7572
16.7501 | -35.0381 | -27.7652
16.7562 | -35.0387 | -27.7733
16.7623 | -35.0392 | -27.7811
16.7684 | -35.0398 | -27.7892
16.7745 | -35.0404 | -27.7973
16.7806 | -35.041 | -27.805
16.7867 | -35.0416 | -27.813
16.7928 | -35.0422 | -27.8209
16.7989 | -35.0427 | -27.8289
16.805 | -35.0433 | -27.8369
16.8111 | -35.0439 | -27.8448
16.8172 | -35.0445 | -27.8528
16.8233 | -35.0451 | -27.8608
16.8294 | -35.0457 | -27.8688
16.8356 | -35.0462 | -27.8769
16.8417 | -35.0468 | -27.8849
16.8478 | -35.0474 | -27.8927
16.8539 | -35.048 | -27.9006
16.86 | -35.0486 | -27.9087
16.8661 | -35.0492 | -27.9167
16.8722 | -35.0497 | -27.9247
16.8783 | -35.0503 | -27.9329
16.8844 | -35.0509 | -27.9406
16.8905 | -35.0515 | -27.9485
16.8966 | -35.0521 | -27.9563
16.9027 | -35.0527 | -27.9643
16.9088 | -35.0532 | -27.9722
16.9149 | -35.0538 | -27.9803
16.921 | -35.0544 | -27.988
16.9271 | -35.055 | -27.9961
16.9332 | -35.0556 | -28.0042
16.9393 | -35.0561 | -28.0121
16.9454 | -35.0567 | -28.02
16.9515 | -35.0573 | -28.028
16.9576 | -35.0579 | -28.0358
16.9637 | -35.0585 | -28.0436
16.9698 | -35.0591 | -28.0514
16.9759 | -35.0596 | -28.0593
16.9821 | -35.0602 | -28.0673
16.9882 | -35.0608 | -28.075
16.9943 | -35.0614 | -28.0828
17.0004 | -35.062 | -28.0908
17.0065 | -35.0626 | -28.0987
17.0126 | -35.0632 | -28.1067
17.0187 | -35.0637 | -28.1147
17.0248 | -35.0643 | -28.1226
17.0309 | -35.0649 | -28.1304
17.037 | -35.0655 | -28.1383
17.0431 | -35.0661 | -28.1461
17.0492 | -35.0667 | -28.1541
17.0553 | -35.0672 | -28.1621
17.0614 | -35.0678 | -28.17
17.0675 | -35.0684 | -28.1779
17.0736 | -35.069 | -28.186
17.0797 | -35.0695 | -28.1938
17.0858 | -35.0701 | -28.2019
17.0919 | -35.0707 | -28.2098
17.098 | -35.0713 | -28.2175
17.1041 | -35.0719 | -28.2253
17.1102 | -35.0725 | -28.2332
17.1163 | -35.0731 | -28.241
17.1225 | -35.0736 | -28.2488
17.1286 | -35.0742 | -28.2564
17.1347 | -35.0748 | -28.2644
17.1408 | -35.0754 | -28.2723
17.1469 | -35.0759 | -28.2801
17.153 | -35.0765 | -28.2879
17.1591 | -35.0771 | -28.2958
17.1652 | -35.0777 | -28.3034
17.1713 | -35.0783 | -28.3112
17.1774 | -35.0788 | -28.3192
17.1835 | -35.0794 | -28.327
17.1896 | -35.08 | -28.3349
17.1957 | -35.0806 | -28.3429
17.2018 | -35.0812 | -28.3504
17.2079 | -35.0818 | -28.3583
17.214 | -35.0823 | -28.3662
17.2201 | -35.0829 | -28.374
17.2262 | -35.0835 | -28.382
17.2323 | -35.0841 | -28.3899
17.2384 | -35.0847 | -28.3977
17.2445 | -35.0853 | -28.4056
17.2506 | -35.0859 | -28.4135
17.2567 | -35.0864 | -28.4214
17.2628 | -35.087 | -28.4293
17.269 | -35.0876 | -28.4372
17.2751 | -35.0882 | -28.4451
17.2812 | -35.0888 | -28.4531
17.2873 | -35.0893 | -28.4611
17.2934 | -35.0899 | -28.469
17.2995 | -35.0905 | -28.4769
17.3056 | -35.0911 | -28.4846
17.3117 | -35.0917 | -28.4925
17.3178 | -35.0922 | -28.5004
17.3239 | -35.0928 | -28.5082
17.33 | -35.0934 | -28.516
17.3361 | -35.094 | -28.5239
17.3422 | -35.0945 | -28.5317
17.3483 | -35.0951 | -28.5395
17.3544 | -35.0957 | -28.5474
17.3605 | -35.0962 | -28.5552
17.3666 | -35.0968 | -28.5631
17.3727 | -35.0974 | -28.5708
17.3788 | -35.098 | -28.5785
17.3849 | -35.0986 | -28.5862
17.391 | -35.0991 | -28.5942
17.3971 | -35.0997 | -28.6019
17.4032 | -35.1003 | -28.6098
17.4094 | -35.1009 | -28.6174
17.4155 | -35.1014 | -28.6251
17.4216 | -35.102 | -28.6329
17.4277 | -35.1026 | -28.6408
17.4338 | -35.1032 | -28.6488
17.4399 | -35.1038 | -28.6567
17.446 | -35.1044 | -28.6644
17.4521 | -35.1049 | -28.6722
17.4582 | -35.1055 | -28.6798
17.4643 | -35.1061 | -28.6875
17.4704 | -35.1067 | -28.6954
17.4765 | -35.1073 | -28.7034
17.4826 | -35.1078 | -28.711
17.4887 | -35.1084 | -28.7189
17.4948 | -35.109 | -28.7268
17.5009 | -35.1095 | -28.7345
17.507 | -35.1101 | -28.7424
17.5131 | -35.1107 | -28.7502
17.5192 | -35.1112 | -28.7576
17.5253 | -35.1118 | -28.7655
17.5314 | -35.1123 | -28.7734
17.5375 | -35.1129 | -28.7812
17.5436 | -35.1135 | -28.789
17.5497 | -35.1141 | -28.7969
17.5559 | -35.1146 | -28.8044
17.562 | -35.1152 | -28.8121
17.5681 | -35.1158 | -28.8198
17.5742 | -35.1164 | -28.8275
17.5803 | -35.1169 | -28.8354
17.5864 | -35.1175 | -28.8431
17.5925 | -35.1181 | -28.8509
17.5986 | -35.1187 | -28.8588
17.6047 | -35.1193 | -28.8666
17.6108 | -35.1198 | -28.8743
17.6169 | -35.1204 | -28.882
17.623 | -35.121 | -28.8898
17.6291 | -35.1215 | -28.8975
17.6352 | -35.1221 | -28.9052
17.6413 | -35.1226 | -28.913
17.6474 | -35.1232 | -28.9209
17.6535 | -35.1238 | -28.9285
17.6596 | -35.1244 | -28.9362
17.6657 | -35.125 | -28.9439
17.6718 | -35.1255 | -28.9517
17.6779 | -35.1261 | -28.9594
17.684 | -35.1267 | -28.9672
17.6901 | -35.1273 | -28.975
17.6963 | -35.1279 | -28.9825
17.7024 | -35.1284 | -28.9902
17.7085 | -35.129 | -28.9979
17.7146 | -35.1296 | -29.0056
17.7207 | -35.1302 | -29.0133
17.7268 | -35.1308 | -29.0209
17.7329 | -35.1314 | -29.0288
17.739 | -35.132 | -29.0368
17.7451 | -35.1325 | -29.0445
17.7512 | -35.1331 | -29.0524
17.7573 | -35.1337 | -29.0602
17.7634 | -35.1342 | -29.0676
17.7695 | -35.1348 | -29.0754
17.7756 | -35.1354 | -29.0831
17.7817 | -35.136 | -29.0909
17.7878 | -35.1365 | -29.0987
17.7939 | -35.1371 | -29.1066
17.8 | -35.1377 | -29.1145
17.8061 | -35.1382 | -29.1222
17.8122 | -35.1388 | -29.13
17.8183 | -35.1394 | -29.1377
17.8244 | -35.14 | -29.1453
17.8305 | -35.1405 | -29.153
17.8366 | -35.1411 | -29.1604
17.8428 | -35.1417 | -29.1682
17.8489 | -35.1422 | -29.176
17.855 | -35.1428 | -29.1837
17.8611 | -35.1433 | -29.1915
17.8672 | -35.1439 | -29.199
17.8733 | -35.1445 | -29.2069
17.8794 | -35.1451 | -29.2145
17.8855 | -35.1457 | -29.2222
17.8916 | -35.1462 | -29.23
17.8977 | -35.1468 | -29.2378
17.9038 | -35.1474 | -29.2452
17.9099 | -35.1479 | -29.2528
17.916 | -35.1485 | -29.2607
17.9221 | -35.1491 | -29.2684
17.9282 | -35.1496 | -29.276
17.9343 | -35.1502 | -29.2839
17.9404 | -35.1507 | -29.2913
17.9465 | -35.1513 | -29.299
17.9526 | -35.1519 | -29.3068
17.9587 | -35.1525 | -29.3145
17.9648 | -35.153 | -29.3222
17.9709 | -35.1536 | -29.3299
17.977 | -35.1541 | -29.3375
17.9832 | -35.1547 | -29.3453
17.9893 | -35.1553 | -29.353
17.9954 | -35.1559 | -29.3607
18.0015 | -35.1564 | -29.3685
18.0076 | -35.157 | -29.3759
18.0137 | -35.1576 | -29.3837
18.0198 | -35.1581 | -29.3913
18.0259 | -35.1587 | -29.3991
18.032 | -35.1593 | -29.4068
18.0381 | -35.1599 | -29.4144
18.0442 | -35.1604 | -29.422
18.0503 | -35.161 | -29.4299
18.0564 | -35.1615 | -29.4375
18.0625 | -35.1621 | -29.4454
18.0686 | -35.1627 | -29.4531
18.0747 | -35.1633 | -29.4606
18.0808 | -35.1638 | -29.468
18.0869 | -35.1644 | -29.4758
18.093 | -35.165 | -29.4836
18.0991 | -35.1655 | -29.4914
18.1052 | -35.1661 | -29.4991
18.1113 | -35.1667 | -29.5067
18.1174 | -35.1672 | -29.5144
18.1236 | -35.1678 | -29.5221
18.1297 | -35.1684 | -29.5299
18.1358 | -35.169 | -29.5375
18.1419 | -35.1695 | -29.5452
18.148 | -35.1701 | -29.5528
18.1541 | -35.1707 | -29.5603
18.1602 | -35.1713 | -29.5681
18.1663 | -35.1719 | -29.5757
18.1724 | -35.1724 | -29.5834
18.1785 | -35.173 | -29.5911
18.1846 | -35.1736 | -29.5985
18.1907 | -35.1741 | -29.6061
18.1968 | -35.1747 | -29.6137
18.2029 | -35.1753 | -29.6213
18.209 | -35.1758 | -29.6291
18.2151 | -35.1764 | -29.6369
18.2212 | -35.177 | -29.6442
18.2273 | -35.1776 | -29.6519
18.2334 | -35.1781 | -29.6596
18.2395 | -35.1787 | -29.6673
18.2456 | -35.1793 | -29.6748
18.2517 | -35.1798 | -29.6824
18.2578 | -35.1804 | -29.6899
18.2639 | -35.181 | -29.6978
18.2701 | -35.1815 | -29.7054
18.2762 | -35.1821 | -29.7129
18.2823 | -35.1827 | -29.7207
18.2884 | -35.1832 | -29.7284
18.2945 | -35.1838 | -29.7358
18.3006 | -35.1844 | -29.7434
18.3067 | -35.1849 | -29.7511
18.3128 | -35.1855 | -29.7588
18.3189 | -35.1861 | -29.7666
18.325 | -35.1866 | -29.7742
18.3311 | -35.1872 | -29.7819
18.3372 | -35.1878 | -29.7895
18.3433 | -35.1883 | -29.7974
18.3494 | -35.1889 | -29.8049
18.3555 | -35.1895 | -29.8125
18.3616 | -35.1901 | -29.82
18.3677 | -35.1907 | -29.8276
18.3738 | -35.1912 | -29.8351
18.3799 | -35.1918 | -29.8428
18.386 | -35.1924 | -29.8504
18.3921 | -35.1929 | -29.8581
18.3982 | -35.1935 | -29.8656
18.4043 | -35.194 | -29.8733
18.4105 | -35.1946 | -29.881
18.4166 | -35.1951 | -29.8885
18.4227 | -35.1957 | -29.8961
18.4288 | -35.1962 | -29.9038
18.4349 | -35.1968 | -29.9115
18.441 | -35.1973 | -29.919
18.4471 | -35.1979 | -29.9264
18.4532 | -35.1984 | -29.9342
18.4593 | -35.199 | -29.9419
18.4654 | -35.1995 | -29.9491
18.4715 | -35.2001 | -29.9569
18.4776 | -35.2007 | -29.9647
18.4837 | -35.2012 | -29.9726
18.4898 | -35.2018 | -29.9802
18.4959 | -35.2024 | -29.9878
18.502 | -35.2029 | -29.9954
18.5081 | -35.2035 | -30.0032
18.5142 | -35.204 | -30.0108
18.5203 | -35.2046 | -30.0183
18.5264 | -35.2052 | -30.0258
18.5325 | -35.2058 | -30.0334
18.5386 | -35.2063 | -30.0409
18.5447 | -35.2069 | -30.0484
18.5508 | -35.2075 | -30.0559
18.557 | -35.208 | -30.0637
18.5631 | -35.2086 | -30.0712
18.5692 | -35.2091 | -30.0789
18.5753 | -35.2097 | -30.0864
18.5814 | -35.2103 | -30.0939
18.5875 | -35.2108 | -30.1012
18.5936 | -35.2114 | -30.1089
18.5997 | -35.2119 | -30.1165
18.6058 | -35.2125 | -30.1239
18.6119 | -35.213 | -30.1315
18.618 | -35.2136 | -30.1392
18.6241 | -35.2142 | -30.1467
18.6302 | -35.2147 | -30.1545
18.6363 | -35.2153 | -30.1622
18.6424 | -35.2158 | -30.1697
18.6485 | -35.2164 | -30.1773
18.6546 | -35.2169 | -30.1849
18.6607 | -35.2175 | -30.1925
18.6668 | -35.2181 | -30.2001
18.6729 | -35.2187 | -30.2078
18.679 | -35.2192 | -30.2152
18.6851 | -35.2198 | -30.2227
18.6912 | -35.2204 | -30.2304
18.6974 | -35.2209 | -30.2381
18.7035 | -35.2215 | -30.2458
18.7096 | -35.222 | -30.2534
18.7157 | -35.2226 | -30.2609
18.7218 | -35.2232 | -30.2685
18.7279 | -35.2237 | -30.2762
18.734 | -35.2243 | -30.2839
18.7401 | -35.2249 | -30.2915
18.7462 | -35.2255 | -30.2989
18.7523 | -35.2261 | -30.3064
18.7584 | -35.2266 | -30.314
18.7645 | -35.2272 | -30.3213
18.7706 | -35.2277 | -30.3289
18.7767 | -35.2283 | -30.3367
18.7828 | -35.2289 | -30.3441
18.7889 | -35.2294 | -30.3514
18.795 | -35.23 | -30.3591
18.8011 | -35.2305 | -30.3667
18.8072 | -35.2311 | -30.3745
18.8133 | -35.2316 | -30.3821
18.8194 | -35.2322 | -30.3894
18.8255 | -35.2328 | -30.397
18.8316 | -35.2333 | -30.4045
18.8377 | -35.2339 | -30.4118
18.8439 | -35.2344 | -30.4194
18.85 | -35.235 | -30.427
18.8561 | -35.2355 | -30.4345
18.8622 | -35.2361 | -30.4421
18.8683 | -35.2367 | -30.4496
18.8744 | -35.2373 | -30.4572
18.8805 | -35.2378 | -30.4649
18.8866 | -35.2384 | -30.4725
18.8927 | -35.2389 | -30.48
18.8988 | -35.2395 | -30.4875
18.9049 | -35.2401 | -30.4948
18.911 | -35.2406 | -30.5025
18.9171 | -35.2412 | -30.5099
18.9232 | -35.2417 | -30.5173
18.9293 | -35.2423 | -30.5248
18.9354 | -35.2428 | -30.5323
18.9415 | -35.2434 | -30.5399
18.9476 | -35.2439 | -30.5474
18.9537 | -35.2445 | -30.5551
18.9598 | -35.245 | -30.5625
18.9659 | -35.2456 | -30.57
18.972 | -35.2462 | -30.5776
18.9781 | -35.2467 | -30.5851
18.9843 | -35.2473 | -30.5927
18.9904 | -35.2478 | -30.6002
18.9965 | -35.2484 | -30.6075
19.0026 | -35.2489 | -30.615
19.0087 | -35.2495 | -30.6225
19.0148 | -35.25 | -30.63
19.0209 | -35.2506 | -30.6375
19.027 | -35.2511 | -30.6451
19.0331 | -35.2517 | -30.6523
19.0392 | -35.2523 | -30.6599
19.0453 | -35.2528 | -30.6674
19.0514 | -35.2534 | -30.6749
19.0575 | -35.2539 | -30.6825
19.0636 | -35.2545 | -30.6896
19.0697 | -35.255 | -30.6971
19.0758 | -35.2556 | -30.7045
19.0819 | -35.2561 | -30.7121
19.088 | -35.2567 | -30.7196
19.0941 | -35.2573 | -30.7272
19.1002 | -35.2578 | -30.7346
19.1063 | -35.2583 | -30.742
19.1124 | -35.2589 | -30.7494
19.1185 | -35.2594 | -30.7567
19.1246 | -35.26 | -30.7643
19.1308 | -35.2605 | -30.7719
19.1369 | -35.2611 | -30.7791
19.143 | -35.2617 | -30.7866
19.1491 | -35.2622 | -30.7941
19.1552 | -35.2628 | -30.8015
19.1613 | -35.2633 | -30.809
19.1674 | -35.2639 | -30.8165
19.1735 | -35.2644 | -30.8238
19.1796 | -35.265 | -30.8312
19.1857 | -35.2655 | -30.8387
19.1918 | -35.2661 | -30.8462
19.1979 | -35.2667 | -30.8537
19.204 | -35.2672 | -30.8609
19.2101 | -35.2678 | -30.8683
19.2162 | -35.2683 | -30.8757
19.2223 | -35.2689 | -30.8832
19.2284 | -35.2694 | -30.8907
19.2345 | -35.2699 | -30.8983
19.2406 | -35.2705 | -30.9055
19.2467 | -35.271 | -30.913
19.2528 | -35.2716 | -30.9205
19.2589 | -35.2722 | -30.9279
19.265 | -35.2727 | -30.9354
19.2712 | -35.2733 | -30.9427
19.2773 | -35.2738 | -30.9499
19.2834 | -35.2744 | -30.9574
19.2895 | -35.2749 | -30.9649
19.2956 | -35.2755 | -30.9724
19.3017 | -35.276 | -30.9799
19.3078 | -35.2766 | -30.9873
19.3139 | -35.2771 | -30.9945
19.32 | -35.2777 | -31.0019
19.3261 | -35.2782 | -31.0093
19.3322 | -35.2787 | -31.0168
19.3383 | -35.2793 | -31.024
19.3444 | -35.2798 | -31.0313
19.3505 | -35.2804 | -31.0387
19.3566 | -35.2809 | -31.0463
19.3627 | -35.2815 | -31.0536
19.3688 | -35.282 | -31.061
19.3749 | -35.2826 | -31.0685
19.381 | -35.2831 | -31.0757
19.3871 | -35.2837 | -31.083
19.3932 | -35.2842 | -31.0905
19.3993 | -35.2848 | -31.0979
19.4054 | -35.2853 | -31.1053
19.4115 | -35.2859 | -31.1128
19.4177 | -35.2864 | -31.12
19.4238 | -35.287 | -31.1273
19.4299 | -35.2875 | -31.1345
19.436 | -35.2881 | -31.142
19.4421 | -35.2886 | -31.1493
19.4482 | -35.2892 | -31.1567
19.4543 | -35.2897 | -31.1637
19.4604 | -35.2902 | -31.1712
19.4665 | -35.2908 | -31.1786
19.4726 | -35.2913 | -31.1857
19.4787 | -35.2919 | -31.1932
19.4848 | -35.2924 | -31.2006
19.4909 | -35.2929 | -31.2078
19.497 | -35.2935 | -31.2152
19.5031 | -35.294 | -31.2224
19.5092 | -35.2946 | -31.2299
19.5153 | -35.2951 | -31.2373
19.5214 | -35.2957 | -31.2445
19.5275 | -35.2962 | -31.2519
19.5336 | -35.2968 | -31.2591
19.5397 | -35.2973 | -31.2664
19.5458 | -35.2979 | -31.2738
19.5519 | -35.2984 | -31.2812
19.5581 | -35.2989 | -31.2884
19.5642 | -35.2995 | -31.2957
19.5703 | -35.3 | -31.3031
19.5764 | -35.3005 | -31.3105
19.5825 | -35.3011 | -31.3178
19.5886 | -35.3017 | -31.3252
19.5947 | -35.3022 | -31.3321
19.6008 | -35.3027 | -31.3396
19.6069 | -35.3033 | -31.3467
19.613 | -35.3038 | -31.3542
19.6191 | -35.3044 | -31.3615
19.6252 | -35.3049 | -31.3688
19.6313 | -35.3055 | -31.3756
19.6374 | -35.306 | -31.3831
19.6435 | -35.3065 | -31.3902
19.6496 | -35.3071 | -31.3974
19.6557 | -35.3076 | -31.4049
19.6618 | -35.3082 | -31.4119
19.6679 | -35.3087 | -31.4194
19.674 | -35.3093 | -31.4268
19.6801 | -35.3098 | -31.4338
19.6862 | -35.3104 | -31.4412
19.6923 | -35.3109 | -31.4483
19.6984 | -35.3115 | -31.4556
19.7046 | -35.312 | -31.4628
19.7107 | -35.3126 | -31.4699
19.7168 | -35.3131 | -31.477
19.7229 | -35.3137 | -31.4844
19.729 | -35.3142 | -31.4918
19.7351 | -35.3148 | -31.4988
19.7412 | -35.3153 | -31.5062
19.7473 | -35.3158 | -31.5136
19.7534 | -35.3163 | -31.5211
19.7595 | -35.3169 | -31.5283
19.7656 | -35.3174 | -31.5357
19.7717 | -35.3179 | -31.5425
19.7778 | -35.3185 | -31.5496
19.7839 | -35.319 | -31.5569
19.79 | -35.3195 | -31.5643
19.7961 | -35.32 | -31.5717
19.8022 | -35.3206 | -31.5787
19.8083 | -35.3211 | -31.5858
19.8144 | -35.3217 | -31.5931
19.8205 | -35.3222 | -31.6003
19.8266 | -35.3227 | -31.6076
19.8327 | -35.3232 | -31.6149
19.8388 | -35.3237 | -31.6218
19.845 | -35.3242 | -31.6292
19.8511 | -35.3248 | -31.6365
19.8572 | -35.3253 | -31.6437
19.8633 | -35.3258 | -31.651
19.8694 | -35.3264 | -31.6583
19.8755 | -35.3269 | -31.6654
19.8816 | -35.3275 | -31.6729
19.8877 | -35.328 | -31.6801
19.8938 | -35.3285 | -31.6875
19.8999 | -35.329 | -31.6948
19.906 | -35.3295 | -31.702
19.9121 | -35.3301 | -31.709
19.9182 | -35.3306 | -31.7164
19.9243 | -35.3311 | -31.7236
19.9304 | -35.3317 | -31.7309
19.9365 | -35.3322 | -31.7384
19.9426 | -35.3327 | -31.7455
19.9487 | -35.3333 | -31.7528
19.9548 | -35.3338 | -31.76
19.9609 | -35.3344 | -31.7671
19.967 | -35.3349 | -31.7742
19.9731 | -35.3355 | -31.7813
19.9792 | -35.336 | -31.7881
19.9853 | -35.3366 | -31.7952
19.9915 | -35.3372 | -31.8023
19.9976 | -35.3377 | -31.8094
20.0037 | -35.3382 | -31.8167
20.0098 | -35.3387 | -31.8242
20.0159 | -35.3392 | -31.8313
20.022 | -35.3397 | -31.8385
20.0281 | -35.3403 | -31.8456
20.0342 | -35.3408 | -31.8527
20.0403 | -35.3413 | -31.8599
20.0464 | -35.3419 | -31.8671
20.0525 | -35.3425 | -31.874
20.0586 | -35.343 | -31.8812
20.0647 | -35.3435 | -31.8884
20.0708 | -35.3441 | -31.8957
20.0769 | -35.3446 | -31.9031
20.083 | -35.3451 | -31.9103
20.0891 | -35.3457 | -31.9174
20.0952 | -35.3462 | -31.9247
20.1013 | -35.3467 | -31.9318
20.1074 | -35.3473 | -31.9389
20.1135 | -35.3478 | -31.9459
20.1196 | -35.3483 | -31.9525
20.1257 | -35.3488 | -31.9597
20.1319 | -35.3494 | -31.9668
20.138 | -35.3499 | -31.9741
20.1441 | -35.3505 | -31.9811
20.1502 | -35.351 | -31.9883
20.1563 | -35.3515 | -31.9949
20.1624 | -35.352 | -32.0018
20.1685 | -35.3525 | -32.0089
20.1746 | -35.353 | -32.016
20.1807 | -35.3535 | -32.0231
20.1868 | -35.3541 | -32.0303
20.1929 | -35.3546 | -32.0371
20.199 | -35.3551 | -32.0443
20.2051 | -35.3556 | -32.0513
20.2112 | -35.3562 | -32.0584
20.2173 | -35.3568 | -32.0652
20.2234 | -35.3573 | -32.0727
20.2295 | -35.3578 | -32.0797
20.2356 | -35.3583 | -32.0871
20.2417 | -35.3588 | -32.0945
20.2478 | -35.3594 | -32.1013
20.2539 | -35.3599 | -32.1082
20.26 | -35.3604 | -32.1146
20.2661 | -35.361 | -32.1217
20.2723 | -35.3615 | -32.1287
20.2784 | -35.362 | -32.1359
20.2845 | -35.3625 | -32.1429
20.2906 | -35.363 | -32.15
20.2967 | -35.3635 | -32.1568
20.3028 | -35.364 | -32.1638
20.3089 | -35.3646 | -32.1707
20.315 | -35.3651 | -32.1779
20.3211 | -35.3657 | -32.1848
20.3272 | -35.3662 | -32.1919
20.3333 | -35.3667 | -32.1988
20.3394 | -35.3673 | -32.2059
20.3455 | -35.3678 | -32.2131
20.3516 | -35.3684 | -32.2201
20.3577 | -35.3689 | -32.2271
20.3638 | -35.3694 | -32.2342
20.3699 | -35.37 | -32.2407
20.376 | -35.3705 | -32.2475
20.3821 | -35.371 | -32.2548
20.3882 | -35.3715 | -32.2617
20.3943 | -35.3721 | -32.2687
20.4004 | -35.3726 | -32.2756
20.4065 | -35.3731 | -32.2829
20.4126 | -35.3736 | -32.2899
20.4188 | -35.3741 | -32.2969
20.4249 | -35.3746 | -32.3038
20.431 | -35.3752 | -32.3107
20.4371 | -35.3757 | -32.3173
20.4432 | -35.3762 | -32.3242
20.4493 | -35.3767 | -32.3312
20.4554 | -35.3773 | -32.3383
20.4615 | -35.3778 | -32.3452
20.4676 | -35.3783 | -32.3522
20.4737 | -35.3789 | -32.359
20.4798 | -35.3794 | -32.3659
20.4859 | -35.3799 | -32.3731
20.492 | -35.3805 | -32.3799
20.4981 | -35.3811 | -32.387
20.5042 | -35.3815 | -32.3942
20.5103 | -35.382 | -32.4006
20.5164 | -35.3825 | -32.4074
20.5225 | -35.383 | -32.4142
20.5286 | -35.3835 | -32.4211
20.5347 | -35.3841 | -32.4282
20.5408 | -35.3845 | -32.4349
20.5469 | -35.385 | -32.4418
20.553 | -35.3855 | -32.449
20.5592 | -35.386 | -32.4559
20.5653 | -35.3865 | -32.4627
20.5714 | -35.387 | -32.4699
20.5775 | -35.3875 | -32.4767
20.5836 | -35.388 | -32.4838
20.5897 | -35.3885 | -32.4908
20.5958 | -35.389 | -32.4981
20.6019 | -35.3895 | -32.5051
20.608 | -35.3901 | -32.5125
20.6141 | -35.3906 | -32.5195
20.6202 | -35.3911 | -32.5265
20.6263 | -35.3917 | -32.5331
20.6324 | -35.3922 | -32.5399
20.6385 | -35.3927 | -32.5468
20.6446 | -35.3932 | -32.5538
20.6507 | -35.3937 | -32.5608
20.6568 | -35.3943 | -32.5681
20.6629 | -35.3948 | -32.575
20.669 | -35.3953 | -32.582
20.6751 | -35.3958 | -32.5891
20.6812 | -35.3964 | -32.5957
20.6873 | -35.3969 | -32.6025
20.6934 | -35.3974 | -32.6095
20.6995 | -35.398 | -32.6166
20.7057 | -35.3985 | -32.6235
20.7118 | -35.399 | -32.6304
20.7179 | -35.3995 | -32.637
20.724 | -35.4001 | -32.644
20.7301 | -35.4006 | -32.651
20.7362 | -35.4012 | -32.6582
20.7423 | -35.4017 | -32.6652
20.7484 | -35.4023 | -32.6722
20.7545 | -35.4028 | -32.6785
20.7606 | -35.4034 | -32.6854
20.7667 | -35.4039 | -32.6923
20.7728 | -35.4044 | -32.6994
20.7789 | -35.405 | -32.7061
20.785 | -35.4055 | -32.7132
20.7911 | -35.406 | -32.7195
20.7972 | -35.4065 | -32.7262
20.8033 | -35.4071 | -32.7329
20.8094 | -35.4076 | -32.7394
20.8155 | -35.408 | -32.7459
20.8216 | -35.4086 | -32.7532
20.8277 | -35.409 | -32.76
20.8338 | -35.4096 | -32.7668
20.8399 | -35.4101 | -32.7734
20.8461 | -35.4106 | -32.7802
20.8522 | -35.4112 | -32.7871
20.8583 | -35.4118 | -32.7939
20.8644 | -35.4123 | -32.8011
20.8705 | -35.4128 | -32.8082
20.8766 | -35.4133 | -32.8149
20.8827 | -35.4138 | -32.822
20.8888 | -35.4143 | -32.8294
20.8949 | -35.4148 | -32.8359
20.901 | -35.4154 | -32.843
20.9071 | -35.4159 | -32.8497
20.9132 | -35.4164 | -32.8565
20.9193 | -35.4169 | -32.8633
20.9254 | -35.4175 | -32.8703
20.9315 | -35.418 | -32.8769
20.9376 | -35.4185 | -32.8837
20.9437 | -35.419 | -32.8907
20.9498 | -35.4195 | -32.8977
20.9559 | -35.42 | -32.905
20.962 | -35.4205 | -32.912
20.9681 | -35.421 | -32.9178
20.9742 | -35.4215 | -32.9244
20.9803 | -35.4221 | -32.9309
20.9864 | -35.4226 | -32.9378
20.9926 | -35.4232 | -32.9447
20.9987 | -35.4237 | -32.9508
21.0048 | -35.4242 | -32.9574
21.0109 | -35.4247 | -32.964
21.017 | -35.4252 | -32.9709
21.0231 | -35.4258 | -32.9779
21.0292 | -35.4263 | -32.9853
21.0353 | -35.4269 | -32.9923
21.0414 | -35.4274 | -32.999
21.0475 | -35.4279 | -33.0058
21.0536 | -35.4285 | -33.0128
21.0597 | -35.429 | -33.0197
21.0658 | -35.4296 | -33.0263
21.0719 | -35.4301 | -33.0325
21.078 | -35.4306 | -33.0396
21.0841 | -35.4312 | -33.0463
21.0902 | -35.4318 | -33.053
21.0963 | -35.4322 | -33.0601
21.1024 | -35.4327 | -33.067
21.1085 | -35.4332 | -33.0732
21.1146 | -35.4337 | -33.0801
21.1207 | -35.4342 | -33.0869
21.1268 | -35.4348 | -33.0934
21.133 | -35.4353 | -33.1
21.1391 | -35.4357 | -33.1059
21.1452 | -35.4362 | -33.1128
21.1513 | -35.4367 | -33.1196
21.1574 | -35.4373 | -33.126
21.1635 | -35.4379 | -33.133
21.1696 | -35.4384 | -33.14
21.1757 | -35.4389 | -33.1467
21.1818 | -35.4394 | -33.1534
21.1879 | -35.44 | -33.1601
21.194 | -35.4404 | -33.1671
21.2001 | -35.441 | -33.1741
21.2062 | -35.4416 | -33.1804
21.2123 | -35.4421 | -33.1871
21.2184 | -35.4427 | -33.1936
21.2245 | -35.443 | -33.2005
21.2306 | -35.4436 | -33.2073
21.2367 | -35.4443 | -33.2132
21.2428 | -35.4448 | -33.2208
21.2489 | -35.4453 | -33.2277
21.255 | -35.4457 | -33.234
21.2611 | -35.4463 | -33.2408
21.2672 | -35.4469 | -33.2474
21.2733 | -35.4475 | -33.254
21.2795 | -35.448 | -33.2609
21.2856 | -35.4485 | -33.268
21.2917 | -35.4491 | -33.2742
21.2978 | -35.4496 | -33.2811
21.3039 | -35.4502 | -33.2876
21.31 | -35.4507 | -33.2943
21.3161 | -35.4512 | -33.3011
21.3222 | -35.4518 | -33.3081
21.3283 | -35.4523 | -33.3145
21.3344 | -35.4528 | -33.3208
21.3405 | -35.4534 | -33.3279
21.3466 | -35.454 | -33.3346
21.3527 | -35.4546 | -33.3412
21.3588 | -35.4551 | -33.3475
21.3649 | -35.4557 | -33.3546
21.371 | -35.4562 | -33.3616
21.3771 | -35.4567 | -33.3689
21.3832 | -35.4573 | -33.3754
21.3893 | -35.4579 | -33.3815
21.3954 | -35.4585 | -33.3886
21.4015 | -35.459 | -33.3952
21.4076 | -35.4595 | -33.4015
21.4137 | -35.46 | -33.4082
21.4199 | -35.4605 | -33.4147
21.426 | -35.4609 | -33.421
21.4321 | -35.4614 | -33.428
21.4382 | -35.4619 | -33.4349
21.4443 | -35.4624 | -33.4408
21.4504 | -35.463 | -33.4473
21.4565 | -35.4635 | -33.4535
21.4626 | -35.4641 | -33.4602
21.4687 | -35.4646 | -33.4667
21.4748 | -35.4651 | -33.4732
21.4809 | -35.4656 | -33.4802
21.487 | -35.4661 | -33.487
21.4931 | -35.4667 | -33.4936
21.4992 | -35.4672 | -33.5004
21.5053 | -35.4677 | -33.5074
21.5114 | -35.4682 | -33.5139
21.5175 | -35.4688 | -33.5207
21.5236 | -35.4693 | -33.5276
21.5297 | -35.4698 | -33.5344
21.5358 | -35.4704 | -33.5413
21.5419 | -35.4709 | -33.548
21.548 | -35.4714 | -33.5545
21.5541 | -35.4719 | -33.5617
21.5602 | -35.4723 | -33.5678
21.5664 | -35.4728 | -33.574
21.5725 | -35.4734 | -33.5806
21.5786 | -35.4739 | -33.5878
21.5847 | -35.4744 | -33.595
21.5908 | -35.475 | -33.602
21.5969 | -35.4755 | -33.6082
21.603 | -35.4761 | -33.6151
21.6091 | -35.4766 | -33.6222
21.6152 | -35.4771 | -33.6293
21.6213 | -35.4776 | -33.6359
21.6274 | -35.4782 | -33.6431
21.6335 | -35.4787 | -33.6498
21.6396 | -35.4793 | -33.6563
21.6457 | -35.4798 | -33.6629
21.6518 | -35.4804 | -33.6702
21.6579 | -35.4808 | -33.6768
21.664 | -35.4815 | -33.6835
21.6701 | -35.4819 | -33.6897
21.6762 | -35.4825 | -33.6965
21.6823 | -35.483 | -33.7033
21.6884 | -35.4835 | -33.71
21.6945 | -35.4841 | -33.7168
21.7006 | -35.4846 | -33.7239
21.7068 | -35.4851 | -33.7295
21.7129 | -35.4857 | -33.7362
21.719 | -35.4862 | -33.7436
21.7251 | -35.4867 | -33.75
21.7312 | -35.4873 | -33.7567
21.7373 | -35.4878 | -33.7625
21.7434 | -35.4883 | -33.7691
21.7495 | -35.4888 | -33.7755
21.7556 | -35.4894 | -33.7823
21.7617 | -35.4899 | -33.7894
21.7678 | -35.4904 | -33.796
21.7739 | -35.4909 | -33.802
21.78 | -35.4914 | -33.8083
21.7861 | -35.4921 | -33.815
21.7922 | -35.4925 | -33.8223
21.7983 | -35.4931 | -33.8284
21.8044 | -35.4937 | -33.8355
21.8105 | -35.4941 | -33.842
21.8166 | -35.4944 | -33.8483
21.8227 | -35.4949 | -33.855
21.8288 | -35.4954 | -33.8619
21.8349 | -35.496 | -33.8688
21.841 | -35.4965 | -33.8751
21.8471 | -35.4969 | -33.8825
21.8533 | -35.4974 | -33.8889
21.8594 | -35.4979 | -33.8952
21.8655 | -35.4985 | -33.9019
21.8716 | -35.499 | -33.9085
21.8777 | -35.4995 | -33.9146
21.8838 | -35.5 | -33.9211
21.8899 | -35.5005 | -33.9279
21.896 | -35.501 | -33.9348
21.9021 | -35.5015 | -33.9416
21.9082 | -35.502 | -33.9479
21.9143 | -35.5026 | -33.9542
21.9204 | -35.5032 | -33.9612
21.9265 | -35.5036 | -33.9679
21.9326 | -35.5041 | -33.9747
21.9387 | -35.5046 | -33.9812
21.9448 | -35.5052 | -33.9877
21.9509 | -35.5058 | -33.9946
21.957 | -35.5063 | -34.0014
21.9631 | -35.5069 | -34.0081
21.9692 | -35.5074 | -34.0151
21.9753 | -35.5079 | -34.0211
21.9814 | -35.5085 | -34.0278
21.9875 | -35.509 | -34.0336
21.9937 | -35.5095 | -34.04
21.9998 | -35.51 | -34.0466
22.0059 | -35.5106 | -34.0531
22.012 | -35.5111 | -34.0595
22.0181 | -35.5116 | -34.0664
22.0242 | -35.5121 | -34.0727
22.0303 | -35.5126 | -34.0792
22.0364 | -35.5132 | -34.0863
22.0425 | -35.5137 | -34.0931
22.0486 | -35.5142 | -34.0995
22.0547 | -35.5147 | -34.1057
22.0608 | -35.5152 | -34.1121
22.0669 | -35.5157 | -34.1185
22.073 | -35.5162 | -34.1249
22.0791 | -35.5168 | -34.1316
22.0852 | -35.5173 | -34.1379
22.0913 | -35.5178 | -34.1442
22.0974 | -35.5183 | -34.151
22.1035 | -35.5188 | -34.1581
22.1096 | -35.5193 | -34.1645
22.1157 | -35.5199 | -34.1711
22.1218 | -35.5204 | -34.1771
22.1279 | -35.5209 | -34.1824
22.134 | -35.5215 | -34.1888
22.1402 | -35.522 | -34.1947
22.1463 | -35.5225 | -34.2011
22.1524 | -35.5229 | -34.2075
22.1585 | -35.5235 | -34.2135
22.1646 | -35.5241 | -34.2197
22.1707 | -35.5246 | -34.2264
22.1768 | -35.5252 | -34.2325
22.1829 | -35.5257 | -34.2393
22.189 | -35.5263 | -34.2459
22.1951 | -35.5268 | -34.2524
22.2012 | -35.5273 | -34.2588
22.2073 | -35.5279 | -34.2655
22.2134 | -35.5284 | -34.272
22.2195 | -35.529 | -34.2782
22.2256 | -35.5295 | -34.2843
22.2317 | -35.53 | -34.2911
22.2378 | -35.5305 | -34.2979
22.2439 | -35.5311 | -34.3046
22.25 | -35.5316 | -34.3111
22.2561 | -35.5321 | -34.3175
22.2622 | -35.5326 | -34.3242
22.2683 | -35.5331 | -34.3304
22.2744 | -35.5336 | -34.3371
22.2806 | -35.5341 | -34.3438
22.2867 | -35.5347 | -34.3506
22.2928 | -35.5353 | -34.3569
22.2989 | -35.5358 | -34.3636
22.305 | -35.5364 | -34.3698
22.3111 | -35.5369 | -34.3764
22.3172 | -35.5375 | -34.3826
22.3233 | -35.538 | -34.389
22.3294 | -35.5385 | -34.3956
22.3355 | -35.539 | -34.4022
22.3416 | -35.5395 | -34.4085
22.3477 | -35.54 | -34.4149
22.3538 | -35.5406 | -34.4216
22.3599 | -35.5411 | -34.4286
22.366 | -35.5416 | -34.435
22.3721 | -35.5422 | -34.4413
22.3782 | -35.5428 | -34.4475
22.3843 | -35.5433 | -34.4542
22.3904 | -35.5439 | -34.4606
22.3965 | -35.5444 | -34.4672
22.4026 | -35.545 | -34.4737
22.4087 | -35.5455 | -34.4796
22.4148 | -35.5459 | -34.486
22.4209 | -35.5464 | -34.4927
22.4271 | -35.5469 | -34.4994
22.4332 | -35.5474 | -34.5054
22.4393 | -35.548 | -34.5117
22.4454 | -35.5485 | -34.5181
22.4515 | -35.549 | -34.5248
22.4576 | -35.5495 | -34.531
22.4637 | -35.55 | -34.5374
22.4698 | -35.5506 | -34.544
22.4759 | -35.5512 | -34.5503
22.482 | -35.5517 | -34.5566
22.4881 | -35.5523 | -34.5636
22.4942 | -35.5528 | -34.5702
22.5003 | -35.5533 | -34.5767
22.5064 | -35.5538 | -34.5831
22.5125 | -35.5544 | -34.589
22.5186 | -35.5549 | -34.5953
22.5247 | -35.5554 | -34.6017
22.5308 | -35.5559 | -34.6081
22.5369 | -35.5565 | -34.6145
22.543 | -35.557 | -34.6208
22.5491 | -35.5574 | -34.6272
22.5552 | -35.558 | -34.6338
22.5613 | -35.5585 | -34.6404
22.5675 | -35.559 | -34.6465
22.5736 | -35.5595 | -34.6529
22.5797 | -35.5601 | -34.6591
22.5858 | -35.5606 | -34.6655
22.5919 | -35.5612 | -34.6719
22.598 | -35.5617 | -34.6782
22.6041 | -35.5622 | -34.6848
22.6102 | -35.5628 | -34.6911
22.6163 | -35.5633 | -34.6975
22.6224 | -35.5639 | -34.7038
22.6285 | -35.5644 | -34.7104
22.6346 | -35.5649 | -34.7167
22.6407 | -35.5655 | -34.7229
22.6468 | -35.566 | -34.7296
22.6529 | -35.5665 | -34.7359
22.659 | -35.5671 | -34.7422
22.6651 | -35.5676 | -34.7486
22.6712 | -35.568 | -34.7551
22.6773 | -35.5686 | -34.7615
22.6834 | -35.5691 | -34.7681
22.6895 | -35.5697 | -34.7746
22.6956 | -35.5702 | -34.7812
22.7017 | -35.5708 | -34.7875
22.7079 | -35.5713 | -34.794
22.714 | -35.5718 | -34.8005
22.7201 | -35.5724 | -34.8069
22.7262 | -35.5729 | -34.813
22.7323 | -35.5735 | -34.8194
22.7384 | -35.574 | -34.8256
22.7445 | -35.5745 | -34.8324
22.7506 | -35.575 | -34.8388
22.7567 | -35.5756 | -34.8454
22.7628 | -35.5761 | -34.8519
22.7689 | -35.5766 | -34.8585
22.775 | -35.5772 | -34.8651
22.7811 | -35.5777 | -34.8716
22.7872 | -35.5782 | -34.878
22.7933 | -35.5787 | -34.8842
22.7994 | -35.5793 | -34.8903
22.8055 | -35.5798 | -34.897
22.8116 | -35.5803 | -34.9036
22.8177 | -35.5809 | -34.9098
22.8238 | -35.5814 | -34.9168
22.8299 | -35.582 | -34.9227
22.836 | -35.5826 | -34.9289
22.8421 | -35.5831 | -34.9354
22.8482 | -35.5837 | -34.942
22.8544 | -35.5842 | -34.9484
22.8605 | -35.5847 | -34.9547
22.8666 | -35.5852 | -34.9609
22.8727 | -35.5857 | -34.9675
22.8788 | -35.5862 | -34.974
22.8849 | -35.5866 | -34.9803
22.891 | -35.5872 | -34.9866
22.8971 | -35.5877 | -34.993
22.9032 | -35.5882 | -34.9996
22.9093 | -35.5887 | -35.006
22.9154 | -35.5892 | -35.0125
22.9215 | -35.5897 | -35.0191
22.9276 | -35.5902 | -35.0254
22.9337 | -35.5907 | -35.0318
22.9398 | -35.5912 | -35.0385
22.9459 | -35.5918 | -35.0448
22.952 | -35.5922 | -35.0508
22.9581 | -35.5927 | -35.0574
22.9642 | -35.5932 | -35.0637
22.9703 | -35.5938 | -35.0696
22.9764 | -35.5943 | -35.0761
22.9825 | -35.5948 | -35.0824
22.9886 | -35.5954 | -35.0886
22.9948 | -35.5959 | -35.0951
23.0009 | -35.5964 | -35.1015
23.007 | -35.597 | -35.1075
23.0131 | -35.5975 | -35.1139
23.0192 | -35.5981 | -35.1205
23.0253 | -35.5986 | -35.127
23.0314 | -35.5991 | -35.1334
23.0375 | -35.5997 | -35.1399
23.0436 | -35.6002 | -35.1461
23.0497 | -35.6007 | -35.1525
23.0558 | -35.6013 | -35.1588
23.0619 | -35.6018 | -35.1652
23.068 | -35.6023 | -35.1714
23.0741 | -35.6029 | -35.1777
23.0802 | -35.6034 | -35.1843
23.0863 | -35.6039 | -35.1905
23.0924 | -35.6044 | -35.197
23.0985 | -35.6049 | -35.2036
23.1046 | -35.6055 | -35.21
23.1107 | -35.606 | -35.2161
23.1168 | -35.6065 | -35.2225
23.1229 | -35.6071 | -35.2289
23.129 | -35.6076 | -35.2353
23.1351 | -35.6081 | -35.2418
23.1413 | -35.6086 | -35.2481
23.1474 | -35.6092 | -35.2546
23.1535 | -35.6097 | -35.2609
23.1596 | -35.6102 | -35.2673
23.1657 | -35.6108 | -35.2735
23.1718 | -35.6113 | -35.2799
23.1779 | -35.6119 | -35.2863
23.184 | -35.6124 | -35.2924
23.1901 | -35.6129 | -35.2987
23.1962 | -35.6134 | -35.3051
23.2023 | -35.614 | -35.3112
23.2084 | -35.6145 | -35.3176
23.2145 | -35.615 | -35.3236
23.2206 | -35.6155 | -35.33
23.2267 | -35.616 | -35.3364
23.2328 | -35.6166 | -35.343
23.2389 | -35.6171 | -35.3494
23.245 | -35.6177 | -35.3556
23.2511 | -35.6182 | -35.3616
23.2572 | -35.6188 | -35.3681
23.2633 | -35.6193 | -35.3745
23.2694 | -35.6198 | -35.3808
23.2755 | -35.6203 | -35.3877
23.2817 | -35.6209 | -35.3941
23.2878 | -35.6214 | -35.4003
23.2939 | -35.622 | -35.4069
23.3 | -35.6225 | -35.4131
23.3061 | -35.6231 | -35.4196
23.3122 | -35.6236 | -35.4262
23.3183 | -35.6241 | -35.4326
23.3244 | -35.6246 | -35.4388
23.3305 | -35.6251 | -35.4453
23.3366 | -35.6257 | -35.4517
23.3427 | -35.6262 | -35.4582
23.3488 | -35.6267 | -35.4648
23.3549 | -35.6272 | -35.471
23.361 | -35.6278 | -35.4771
23.3671 | -35.6284 | -35.4836
23.3732 | -35.6289 | -35.4899
23.3793 | -35.6294 | -35.4962
23.3854 | -35.6299 | -35.5024
23.3915 | -35.6305 | -35.5084
23.3976 | -35.631 | -35.515
23.4037 | -35.6315 | -35.5216
23.4098 | -35.6321 | -35.5279
23.4159 | -35.6326 | -35.5343
23.422 | -35.6331 | -35.5408
23.4282 | -35.6337 | -35.5473
23.4343 | -35.6342 | -35.5533
23.4404 | -35.6347 | -35.5596
23.4465 | -35.6352 | -35.5662
23.4526 | -35.6357 | -35.5724
23.4587 | -35.6362 | -35.5788
23.4648 | -35.6368 | -35.5852
23.4709 | -35.6373 | -35.5916
23.477 | -35.6378 | -35.598
23.4831 | -35.6383 | -35.6044
23.4892 | -35.6389 | -35.6107
23.4953 | -35.6395 | -35.6171
23.5014 | -35.64 | -35.6232
23.5075 | -35.6405 | -35.6293
23.5136 | -35.641 | -35.6358
23.5197 | -35.6416 | -35.6423
23.5258 | -35.6421 | -35.6487
23.5319 | -35.6427 | -35.6548
23.538 | -35.6432 | -35.661
23.5441 | -35.6437 | -35.6675
23.5502 | -35.6442 | -35.6741
23.5563 | -35.6448 | -35.6805
23.5624 | -35.6453 | -35.687
23.5686 | -35.6459 | -35.6932
23.5747 | -35.6464 | -35.6995
23.5808 | -35.6469 | -35.7058
23.5869 | -35.6474 | -35.7121
23.593 | -35.648 | -35.7184
23.5991 | -35.6485 | -35.7244
23.6052 | -35.649 | -35.7306
23.6113 | -35.6495 | -35.737
23.6174 | -35.6501 | -35.7432
23.6235 | -35.6506 | -35.7497
23.6296 | -35.6511 | -35.7559
23.6357 | -35.6516 | -35.7623
23.6418 | -35.6522 | -35.7684
23.6479 | -35.6527 | -35.7749
23.654 | -35.6532 | -35.7809
23.6601 | -35.6538 | -35.7873
23.6662 | -35.6543 | -35.7937
23.6723 | -35.6549 | -35.7999
23.6784 | -35.6554 | -35.8065
23.6845 | -35.6559 | -35.8129
23.6906 | -35.6564 | -35.8192
23.6967 | -35.657 | -35.8252
23.7028 | -35.6575 | -35.8313
23.7089 | -35.658 | -35.8376
23.7151 | -35.6585 | -35.8441
23.7212 | -35.659 | -35.8502
23.7273 | -35.6595 | -35.8565
23.7334 | -35.66 | -35.8629
23.7395 | -35.6606 | -35.8693
23.7456 | -35.6612 | -35.8752
23.7517 | -35.6617 | -35.8816
23.7578 | -35.6622 | -35.8878
23.7639 | -35.6627 | -35.8939
23.77 | -35.6632 | -35.9001
23.7761 | -35.6637 | -35.9067
23.7822 | -35.6642 | -35.9126
23.7883 | -35.6648 | -35.9189
23.7944 | -35.6652 | -35.925
23.8005 | -35.6657 | -35.9315
23.8066 | -35.6663 | -35.9378
23.8127 | -35.6668 | -35.9442
23.8188 | -35.6674 | -35.9506
23.8249 | -35.6679 | -35.9567
23.831 | -35.6684 | -35.9627
23.8371 | -35.6689 | -35.969
23.8432 | -35.6695 | -35.9752
23.8493 | -35.67 | -35.9817
23.8555 | -35.6706 | -35.9882
23.8616 | -35.6711 | -35.9945
23.8677 | -35.6716 | -36.0008
23.8738 | -35.6721 | -36.0072
23.8799 | -35.6726 | -36.0134
23.886 | -35.6732 | -36.0194
23.8921 | -35.6737 | -36.0253
23.8982 | -35.6742 | -36.0313
23.9043 | -35.6747 | -36.0378
23.9104 | -35.6752 | -36.044
23.9165 | -35.6758 | -36.0504
23.9226 | -35.6763 | -36.0563
23.9287 | -35.6768 | -36.0625
23.9348 | -35.6773 | -36.0689
23.9409 | -35.6778 | -36.0754
23.947 | -35.6784 | -36.0816
23.9531 | -35.6789 | -36.0873
23.9592 | -35.6794 | -36.0938
23.9653 | -35.6799 | -36.0998
23.9714 | -35.6804 | -36.106
23.9775 | -35.681 | -36.1119
23.9836 | -35.6815 | -36.118
23.9897 | -35.6821 | -36.1241
23.9958 | -35.6826 | -36.1305
24.002 | -35.6831 | -36.1365
24.0081 | -35.6836 | -36.1428
24.0142 | -35.6842 | -36.1491
24.0203 | -35.6847 | -36.1557
24.0264 | -35.6853 | -36.1619
24.0325 | -35.6858 | -36.1681
24.0386 | -35.6863 | -36.1746
24.0447 | -35.6868 | -36.1803
24.0508 | -35.6873 | -36.1862
24.0569 | -35.6879 | -36.1925
24.063 | -35.6884 | -36.198
24.0691 | -35.6889 | -36.2044
24.0752 | -35.6895 | -36.2108
24.0813 | -35.69 | -36.2169
24.0874 | -35.6905 | -36.2232
24.0935 | -35.691 | -36.2292
24.0996 | -35.6916 | -36.2351
24.1057 | -35.6921 | -36.2416
24.1118 | -35.6926 | -36.2482
24.1179 | -35.6931 | -36.2545
24.124 | -35.6935 | -36.2604
24.1301 | -35.6941 | -36.2665
24.1362 | -35.6946 | -36.2723
24.1424 | -35.6952 | -36.2788
24.1485 | -35.6957 | -36.2851
24.1546 | -35.6963 | -36.2912
24.1607 | -35.6968 | -36.2974
24.1668 | -35.6973 | -36.3036
24.1729 | -35.6978 | -36.3099
24.179 | -35.6983 | -36.3158
24.1851 | -35.6988 | -36.3222
24.1912 | -35.6993 | -36.3283
24.1973 | -35.6998 | -36.3344
24.2034 | -35.7003 | -36.341
24.2095 | -35.7008 | -36.3471
24.2156 | -35.7014 | -36.3534
24.2217 | -35.7019 | -36.3591
24.2278 | -35.7024 | -36.3651
24.2339 | -35.703 | -36.3711
24.24 | -35.7034 | -36.3769
24.2461 | -35.7039 | -36.3835
24.2522 | -35.7044 | -36.3898
24.2583 | -35.705 | -36.3965
24.2644 | -35.7055 | -36.4027
24.2705 | -35.7061 | -36.4092
24.2766 | -35.7066 | -36.4153
24.2827 | -35.7071 | -36.4217
24.2889 | -35.7076 | -36.4281
24.295 | -35.7081 | -36.4341
24.3011 | -35.7085 | -36.4405
24.3072 | -35.709 | -36.4465
24.3133 | -35.7095 | -36.453
24.3194 | -35.7101 | -36.4595
24.3255 | -35.7106 | -36.4657
24.3316 | -35.7111 | -36.4718
24.3377 | -35.7117 | -36.4779
24.3438 | -35.7122 | -36.4839
24.3499 | -35.7127 | -36.4901
24.356 | -35.7132 | -36.4963
24.3621 | -35.7137 | -36.5026
24.3682 | -35.7142 | -36.5089
24.3743 | -35.7147 | -36.5153
24.3804 | -35.7152 | -36.5209
24.3865 | -35.7157 | -36.5271
24.3926 | -35.7162 | -36.5334
24.3987 | -35.7168 | -36.5393
24.4048 | -35.7173 | -36.5454
24.4109 | -35.7178 | -36.5511
24.417 | -35.7183 | -36.5574
24.4231 | -35.7188 | -36.5632
24.4293 | -35.7193 | -36.5692
24.4354 | -35.7198 | -36.5755
24.4415 | -35.7203 | -36.5818
24.4476 | -35.7209 | -36.588
24.4537 | -35.7214 | -36.5941
24.4598 | -35.7219 | -36.6001
24.4659 | -35.7224 | -36.6067
24.472 | -35.7229 | -36.6126
24.4781 | -35.7235 | -36.619
24.4842 | -35.724 | -36.6251
24.4903 | -35.7245 | -36.6312
24.4964 | -35.725 | -36.6375
24.5025 | -35.7255 | -36.6438
24.5086 | -35.726 | -36.6499
24.5147 | -35.7265 | -36.656
24.5208 | -35.727 | -36.6621
24.5269 | -35.7276 | -36.6681
24.533 | -35.7281 | -36.6741
24.5391 | -35.7286 | -36.6802
24.5452 | -35.7291 | -36.6864
24.5513 | -35.7297 | -36.6927
24.5574 | -35.7302 | -36.6986
24.5635 | -35.7307 | -36.7046
24.5696 | -35.7312 | -36.7106
24.5758 | -35.7317 | -36.7169
24.5819 | -35.7323 | -36.7232
24.588 | -35.7328 | -36.7291
24.5941 | -35.7333 | -36.7351
24.6002 | -35.7338 | -36.7408
24.6063 | -35.7344 | -36.7472
24.6124 | -35.735 | -36.7536
24.6185 | -35.7355 | -36.7597
24.6246 | -35.736 | -36.7657
24.6307 | -35.7364 | -36.772
24.6368 | -35.7369 | -36.7782
24.6429 | -35.7374 | -36.7841
24.649 | -35.738 | -36.7904
24.6551 | -35.7385 | -36.7964
24.6612 | -35.739 | -36.8021
24.6673 | -35.7395 | -36.8082
24.6734 | -35.74 | -36.8144
24.6795 | -35.7405 | -36.8203
24.6856 | -35.7411 | -36.8266
24.6917 | -35.7416 | -36.8329
24.6978 | -35.7421 | -36.8392
24.7039 | -35.7426 | -36.8456
24.71 | -35.7431 | -36.8516
24.7162 | -35.7437 | -36.8575
24.7223 | -35.7442 | -36.8639
24.7284 | -35.7447 | -36.8701
24.7345 | -35.7452 | -36.8761
24.7406 | -35.7458 | -36.8825
24.7467 | -35.7463 | -36.8891
24.7528 | -35.7468 | -36.895
24.7589 | -35.7474 | -36.9009
24.765 | -35.7479 | -36.9067
24.7711 | -35.7484 | -36.913
24.7772 | -35.7489 | -36.9192
24.7833 | -35.7494 | -36.9252
24.7894 | -35.7499 | -36.931
24.7955 | -35.7504 | -36.937
24.8016 | -35.7509 | -36.9429
24.8077 | -35.7514 | -36.9494
24.8138 | -35.752 | -36.9555
24.8199 | -35.7525 | -36.9613
24.826 | -35.753 | -36.9673
24.8321 | -35.7536 | -36.9737
24.8382 | -35.7541 | -36.9802
24.8443 | -35.7545 | -36.9859
24.8504 | -35.755 | -36.9917
24.8565 | -35.7556 | -36.9976
24.8627 | -35.7561 | -37.0039
24.8688 | -35.7567 | -37.0104
24.8749 | -35.7572 | -37.0162
24.881 | -35.7578 | -37.0226
24.8871 | -35.7583 | -37.0288
24.8932 | -35.7588 | -37.0351
24.8993 | -35.7594 | -37.0409
24.9054 | -35.7599 | -37.0469
24.9115 | -35.7603 | -37.0533
24.9176 | -35.7609 | -37.0592
24.9237 | -35.7614 | -37.0652
24.9298 | -35.7619 | -37.0713
24.9359 | -35.7624 | -37.0773
24.942 | -35.7629 | -37.0829
24.9481 | -35.7634 | -37.089
24.9542 | -35.7639 | -37.0951
24.9603 | -35.7644 | -37.1012
24.9664 | -35.7649 | -37.1073
24.9725 | -35.7654 | -37.1132
24.9786 | -35.7658 | -37.1186
24.9847 | -35.7664 | -37.1248
24.9908 | -35.7669 | -37.1308
24.9969 | -35.7674 | -37.1371
25.0031 | -35.7679 | -37.1434
25.0092 | -35.7684 | -37.1495
25.0153 | -35.7689 | -37.1553
25.0214 | -35.7695 | -37.1613
25.0275 | -35.77 | -37.167
25.0336 | -35.7705 | -37.1729
25.0397 | -35.771 | -37.1791
25.0458 | -35.7715 | -37.1851
25.0519 | -35.772 | -37.1913
25.058 | -35.7725 | -37.1975
25.0641 | -35.773 | -37.2037
25.0702 | -35.7735 | -37.2099
25.0763 | -35.7739 | -37.2161
25.0824 | -35.7744 | -37.2218
25.0885 | -35.7749 | -37.2282
25.0946 | -35.7754 | -37.2344
25.1007 | -35.776 | -37.2407
25.1068 | -35.7765 | -37.2466
25.1129 | -35.777 | -37.2526
25.119 | -35.7775 | -37.2591
25.1251 | -35.778 | -37.2649
25.1312 | -35.7786 | -37.2712
25.1373 | -35.7791 | -37.2775
25.1435 | -35.7796 | -37.2833
25.1496 | -35.7802 | -37.289
25.1557 | -35.7807 | -37.2948
25.1618 | -35.7812 | -37.3008
25.1679 | -35.7817 | -37.307
25.174 | -35.7821 | -37.3131
25.1801 | -35.7827 | -37.3196
25.1862 | -35.7832 | -37.3255
25.1923 | -35.7837 | -37.3312
25.1984 | -35.7843 | -37.3373
25.2045 | -35.7849 | -37.3433
25.2106 | -35.7854 | -37.3498
25.2167 | -35.7859 | -37.3558
25.2228 | -35.7864 | -37.3616
25.2289 | -35.7869 | -37.3679
25.235 | -35.7875 | -37.374
25.2411 | -35.788 | -37.3803
25.2472 | -35.7885 | -37.3867
25.2533 | -35.789 | -37.3931
25.2594 | -35.7896 | -37.3991
25.2655 | -35.7902 | -37.405
25.2716 | -35.7907 | -37.4105
25.2777 | -35.7912 | -37.4161
25.2838 | -35.7917 | -37.4222
25.29 | -35.7922 | -37.4279
25.2961 | -35.7927 | -37.434
25.3022 | -35.7932 | -37.4403
25.3083 | -35.7938 | -37.4458
25.3144 | -35.7944 | -37.4514
25.3205 | -35.7949 | -37.4579
25.3266 | -35.7953 | -37.4642
25.3327 | -35.7958 | -37.4698
25.3388 | -35.7963 | -37.4759
25.3449 | -35.7969 | -37.4819
25.351 | -35.7974 | -37.488
25.3571 | -35.798 | -37.4939
25.3632 | -35.7986 | -37.4994
25.3693 | -35.7991 | -37.5051
25.3754 | -35.7997 | -37.5113
25.3815 | -35.8002 | -37.5176
25.3876 | -35.8007 | -37.5237
25.3937 | -35.8011 | -37.5298
25.3998 | -35.8017 | -37.5355
25.4059 | -35.8022 | -37.542
25.412 | -35.8027 | -37.5481
25.4181 | -35.8031 | -37.5542
25.4242 | -35.8036 | -37.5602
25.4304 | -35.8041 | -37.5661
25.4365 | -35.8046 | -37.5718
25.4426 | -35.8051 | -37.5777
25.4487 | -35.8057 | -37.5839
25.4548 | -35.8062 | -37.5898
25.4609 | -35.8067 | -37.5954
25.467 | -35.8074 | -37.6013
25.4731 | -35.8079 | -37.6073
25.4792 | -35.8084 | -37.6131
25.4853 | -35.809 | -37.619
25.4914 | -35.8094 | -37.6249
25.4975 | -35.81 | -37.6313
25.5036 | -35.8105 | -37.6375
25.5097 | -35.811 | -37.6432
25.5158 | -35.8115 | -37.6492
25.5219 | -35.812 | -37.6553
25.528 | -35.8125 | -37.6614
25.5341 | -35.813 | -37.6673
25.5402 | -35.8135 | -37.6734
25.5463 | -35.814 | -37.6796
25.5524 | -35.8145 | -37.6853
25.5585 | -35.815 | -37.6912
25.5646 | -35.8155 | -37.6976
25.5707 | -35.816 | -37.7037
25.5769 | -35.8165 | -37.71
25.583 | -35.817 | -37.7162
25.5891 | -35.8175 | -37.7223
25.5952 | -35.818 | -37.7282
25.6013 | -35.8184 | -37.7343
25.6074 | -35.8189 | -37.7408
25.6135 | -35.8195 | -37.7466
25.6196 | -35.8201 | -37.7523
25.6257 | -35.8206 | -37.7583
25.6318 | -35.821 | -37.7644
25.6379 | -35.8215 | -37.7702
25.644 | -35.8219 | -37.776
25.6501 | -35.8224 | -37.7819
25.6562 | -35.823 | -37.7877
25.6623 | -35.8234 | -37.7936
25.6684 | -35.8239 | -37.7996
25.6745 | -35.8244 | -37.8059
25.6806 | -35.8249 | -37.8117
25.6867 | -35.8253 | -37.8178
25.6928 | -35.8258 | -37.8236
25.6989 | -35.8263 | -37.8295
25.705 | -35.8268 | -37.8357
25.7111 | -35.8274 | -37.8413
25.7173 | -35.8279 | -37.8472
25.7234 | -35.8284 | -37.8534
25.7295 | -35.8288 | -37.8591
25.7356 | -35.8293 | -37.8653
25.7417 | -35.8298 | -37.8715
25.7478 | -35.8302 | -37.8771
25.7539 | -35.8308 | -37.883
25.76 | -35.8313 | -37.8888
25.7661 | -35.8318 | -37.8949
25.7722 | -35.8324 | -37.9009
25.7783 | -35.8329 | -37.9069
25.7844 | -35.8334 | -37.9126
25.7905 | -35.8339 | -37.9188
25.7966 | -35.8344 | -37.9248
25.8027 | -35.8349 | -37.9309
25.8088 | -35.8354 | -37.9365
25.8149 | -35.8358 | -37.9426
25.821 | -35.8364 | -37.9485
25.8271 | -35.8369 | -37.9543
25.8332 | -35.8374 | -37.9602
25.8393 | -35.838 | -37.9661
25.8454 | -35.8385 | -37.972
25.8515 | -35.839 | -37.9781
25.8576 | -35.8396 | -37.9835
25.8638 | -35.8401 | -37.9896
25.8699 | -35.8406 | -37.9957
25.876 | -35.8411 | -38.0014
25.8821 | -35.8416 | -38.0075
25.8882 | -35.8421 | -38.0137
25.8943 | -35.8426 | -38.02
25.9004 | -35.843 | -38.0262
25.9065 | -35.8436 | -38.0327
25.9126 | -35.844 | -38.0387
25.9187 | -35.8445 | -38.0446
25.9248 | -35.8451 | -38.0501
25.9309 | -35.8456 | -38.0559
25.937 | -35.8461 | -38.0619
25.9431 | -35.8466 | -38.0675
25.9492 | -35.8471 | -38.0733
25.9553 | -35.8477 | -38.0792
25.9614 | -35.8482 | -38.0852
25.9675 | -35.8487 | -38.0914
25.9736 | -35.8492 | -38.0975
25.9797 | -35.8497 | -38.1034
25.9858 | -35.8503 | -38.1092
25.9919 | -35.8508 | -38.1152
25.998 | -35.8514 | -38.1212
26.0042 | -35.8518 | -38.1271
26.0103 | -35.8523 | -38.133
26.0164 | -35.8527 | -38.1387
26.0225 | -35.8532 | -38.1448
26.0286 | -35.8537 | -38.1506
26.0347 | -35.8541 | -38.1556
26.0408 | -35.8547 | -38.1614
26.0469 | -35.8552 | -38.1667
26.053 | -35.8557 | -38.1729
26.0591 | -35.8562 | -38.1792
26.0652 | -35.8568 | -38.1847
26.0713 | -35.8573 | -38.191
26.0774 | -35.8578 | -38.197
26.0835 | -35.8583 | -38.2027
26.0896 | -35.8588 | -38.2088
26.0957 | -35.8593 | -38.2148
26.1018 | -35.8598 | -38.2204
26.1079 | -35.8604 | -38.2264
26.114 | -35.8609 | -38.2324
26.1201 | -35.8614 | -38.2383
26.1262 | -35.8619 | -38.244
26.1323 | -35.8624 | -38.2501
26.1384 | -35.8628 | -38.2557
26.1445 | -35.8633 | -38.262
26.1507 | -35.8638 | -38.2679
26.1568 | -35.8644 | -38.274
26.1629 | -35.865 | -38.28
26.169 | -35.8655 | -38.2855
26.1751 | -35.8661 | -38.2911
26.1812 | -35.8666 | -38.2969
26.1873 | -35.8671 | -38.3024
26.1934 | -35.8676 | -38.3086
26.1995 | -35.8682 | -38.3142
26.2056 | -35.8687 | -38.3202
26.2117 | -35.8693 | -38.3269
26.2178 | -35.8698 | -38.3326
26.2239 | -35.8703 | -38.3386
26.23 | -35.8708 | -38.3443
26.2361 | -35.8714 | -38.3499
26.2422 | -35.8718 | -38.3557
26.2483 | -35.8723 | -38.3616
26.2544 | -35.8728 | -38.3673
26.2605 | -35.8733 | -38.3731
26.2666 | -35.8738 | -38.3792
26.2727 | -35.8742 | -38.3853
26.2788 | -35.8747 | -38.3911
26.2849 | -35.8752 | -38.397
26.2911 | -35.8757 | -38.4028
26.2972 | -35.8762 | -38.4088
26.3033 | -35.8767 | -38.4145
26.3094 | -35.8772 | -38.4204
26.3155 | -35.8778 | -38.4259
26.3216 | -35.8784 | -38.4315
26.3277 | -35.8789 | -38.4377
26.3338 | -35.8794 | -38.4436
26.3399 | -35.8799 | -38.449
26.346 | -35.8805 | -38.4547
26.3521 | -35.881 | -38.4603
26.3582 | -35.8815 | -38.4659
26.3643 | -35.882 | -38.4723
26.3704 | -35.8825 | -38.4784
26.3765 | -35.8829 | -38.4843
26.3826 | -35.8835 | -38.4899
26.3887 | -35.8841 | -38.4961
26.3948 | -35.8846 | -38.5021
26.4009 | -35.8851 | -38.5083
26.407 | -35.8856 | -38.5145
26.4131 | -35.8861 | -38.5208
26.4192 | -35.8866 | -38.5268
26.4253 | -35.8871 | -38.533
26.4314 | -35.8877 | -38.5391
26.4376 | -35.8882 | -38.5449
26.4437 | -35.8888 | -38.5505
26.4498 | -35.8893 | -38.5563
26.4559 | -35.8899 | -38.5626
26.462 | -35.8904 | -38.5686
26.4681 | -35.891 | -38.5745
26.4742 | -35.8915 | -38.5802
26.4803 | -35.892 | -38.5861
26.4864 | -35.8925 | -38.5913
26.4925 | -35.8929 | -38.5969
26.4986 | -35.8935 | -38.6031
26.5047 | -35.894 | -38.6098
26.5108 | -35.8945 | -38.6154
26.5169 | -35.8949 | -38.6215
26.523 | -35.8955 | -38.6269
26.5291 | -35.896 | -38.6325
26.5352 | -35.8965 | -38.6385
26.5413 | -35.897 | -38.6445
26.5474 | -35.8975 | -38.6506
26.5535 | -35.898 | -38.6565
26.5596 | -35.8985 | -38.6621
26.5657 | -35.899 | -38.6679
26.5718 | -35.8995 | -38.6735
26.578 | -35.9001 | -38.6795
26.5841 | -35.9006 | -38.6855
26.5902 | -35.9011 | -38.6915
26.5963 | -35.9015 | -38.697
26.6024 | -35.9019 | -38.7024
26.6085 | -35.9025 | -38.7082
26.6146 | -35.903 | -38.7139
26.6207 | -35.9036 | -38.7198
26.6268 | -35.904 | -38.7256
26.6329 | -35.9046 | -38.7312
26.639 | -35.9051 | -38.7367
26.6451 | -35.9055 | -38.743
26.6512 | -35.906 | -38.7486
26.6573 | -35.9066 | -38.7545
26.6634 | -35.907 | -38.7607
26.6695 | -35.9076 | -38.7669
26.6756 | -35.9081 | -38.7726
26.6817 | -35.9086 | -38.7787
26.6878 | -35.909 | -38.7848
26.6939 | -35.9095 | -38.7906
26.7 | -35.91 | -38.7963
26.7061 | -35.9105 | -38.8023
26.7122 | -35.911 | -38.8083
26.7183 | -35.9115 | -38.814
26.7245 | -35.912 | -38.82
26.7306 | -35.9125 | -38.8258
26.7367 | -35.9128 | -38.8312
26.7428 | -35.9134 | -38.8372
26.7489 | -35.9139 | -38.8429
26.755 | -35.9144 | -38.8491
26.7611 | -35.9149 | -38.8559
26.7672 | -35.9153 | -38.8614
26.7733 | -35.9159 | -38.8674
26.7794 | -35.9164 | -38.8732
26.7855 | -35.9169 | -38.8791
26.7916 | -35.9174 | -38.8848
26.7977 | -35.9179 | -38.8906
26.8038 | -35.9184 | -38.8965
26.8099 | -35.9189 | -38.9017
26.816 | -35.9194 | -38.9078
26.8221 | -35.9199 | -38.9138
26.8282 | -35.9204 | -38.9191
26.8343 | -35.9209 | -38.9248
26.8404 | -35.9213 | -38.9305
26.8465 | -35.9218 | -38.9361
26.8526 | -35.9224 | -38.9414
26.8587 | -35.9229 | -38.9468
26.8649 | -35.9234 | -38.9527
26.871 | -35.924 | -38.9586
26.8771 | -35.9245 | -38.9641
26.8832 | -35.925 | -38.9697
26.8893 | -35.9256 | -38.9751
26.8954 | -35.9261 | -38.9806
26.9015 | -35.9266 | -38.9859
26.9076 | -35.9272 | -38.9912
26.9137 | -35.9276 | -38.9967
26.9198 | -35.9281 | -39.0029
26.9259 | -35.9286 | -39.0084
26.932 | -35.9291 | -39.0146
26.9381 | -35.9297 | -39.0204
26.9442 | -35.9301 | -39.0255
26.9503 | -35.9306 | -39.0317
26.9564 | -35.9311 | -39.0373
26.9625 | -35.9316 | -39.0433
26.9686 | -35.9321 | -39.0491
26.9747 | -35.9326 | -39.0551
26.9808 | -35.933 | -39.0608
26.9869 | -35.9335 | -39.0664
26.993 | -35.934 | -39.0722
26.9991 | -35.9345 | -39.078
27.0052 | -35.9349 | -39.0838
27.0114 | -35.9355 | -39.0892
27.0175 | -35.9359 | -39.0943
27.0236 | -35.9364 | -39.1004
27.0297 | -35.9369 | -39.1063
27.0358 | -35.9374 | -39.112
27.0419 | -35.9378 | -39.1173
27.048 | -35.9383 | -39.1227
27.0541 | -35.9387 | -39.1283
27.0602 | -35.9392 | -39.1341
27.0663 | -35.9397 | -39.1396
27.0724 | -35.9401 | -39.145
27.0785 | -35.9406 | -39.1513
27.0846 | -35.9411 | -39.1573
27.0907 | -35.9416 | -39.1627
27.0968 | -35.9421 | -39.1685
27.1029 | -35.9426 | -39.1746
27.109 | -35.9431 | -39.1806
27.1151 | -35.9436 | -39.1869
27.1212 | -35.9441 | -39.1923
27.1273 | -35.9446 | -39.1978
27.1334 | -35.9452 | -39.2033
27.1395 | -35.9458 | -39.209
27.1456 | -35.9462 | -39.2146
27.1518 | -35.9467 | -39.2205
27.1579 | -35.947 | -39.2253
27.164 | -35.9476 | -39.2313
27.1701 | -35.948 | -39.2378
27.1762 | -35.9484 | -39.2437
27.1823 | -35.9488 | -39.2498
27.1884 | -35.9494 | -39.2558
27.1945 | -35.9498 | -39.2614
27.2006 | -35.9503 | -39.2671
27.2067 | -35.9507 | -39.273
27.2128 | -35.9512 | -39.2786
27.2189 | -35.9516 | -39.2842
27.225 | -35.9521 | -39.29
27.2311 | -35.9525 | -39.2952
27.2372 | -35.953 | -39.3004
27.2433 | -35.9535 | -39.3059
27.2494 | -35.954 | -39.3118
27.2555 | -35.9545 | -39.3176
27.2616 | -35.955 | -39.3228
27.2677 | -35.9555 | -39.3286
27.2738 | -35.9559 | -39.3341
27.2799 | -35.9564 | -39.3399
27.286 | -35.9569 | -39.346
27.2921 | -35.9573 | -39.3514
27.2983 | -35.9578 | -39.3567
27.3044 | -35.9582 | -39.3627
27.3105 | -35.9588 | -39.3684
27.3166 | -35.9593 | -39.3742
27.3227 | -35.9598 | -39.3801
27.3288 | -35.9603 | -39.3854
27.3349 | -35.9608 | -39.3916
27.341 | -35.9612 | -39.3972
27.3471 | -35.9617 | -39.4024
27.3532 | -35.9621 | -39.4079
27.3593 | -35.9627 | -39.4138
27.3654 | -35.9632 | -39.4192
27.3715 | -35.9637 | -39.4245
27.3776 | -35.9642 | -39.4308
27.3837 | -35.9647 | -39.4364
27.3898 | -35.9652 | -39.4411
27.3959 | -35.9656 | -39.4466
27.402 | -35.9661 | -39.4526
27.4081 | -35.9667 | -39.4587
27.4142 | -35.9671 | -39.4638
27.4203 | -35.9677 | -39.4688
27.4264 | -35.9681 | -39.4748
27.4325 | -35.9687 | -39.4807
27.4387 | -35.9691 | -39.486
27.4448 | -35.9696 | -39.4917
27.4509 | -35.97 | -39.4976
27.457 | -35.9704 | -39.5027
27.4631 | -35.971 | -39.5089
27.4692 | -35.9715 | -39.5143
27.4753 | -35.972 | -39.5204
27.4814 | -35.9725 | -39.5261
27.4875 | -35.973 | -39.5317
27.4936 | -35.9735 | -39.5368
27.4997 | -35.974 | -39.5422
27.5058 | -35.9745 | -39.5479
27.5119 | -35.9751 | -39.5532
27.518 | -35.9755 | -39.5592
27.5241 | -35.976 | -39.5644
27.5302 | -35.9765 | -39.5697
27.5363 | -35.977 | -39.5753
27.5424 | -35.9773 | -39.5809
27.5485 | -35.9778 | -39.5864
27.5546 | -35.9783 | -39.5918
27.5607 | -35.9787 | -39.5978
27.5668 | -35.9793 | -39.6039
27.5729 | -35.9798 | -39.6095
27.5791 | -35.9804 | -39.6152
27.5852 | -35.9809 | -39.6209
27.5913 | -35.9814 | -39.627
27.5974 | -35.9819 | -39.6328
27.6035 | -35.9824 | -39.6379
27.6096 | -35.9829 | -39.6436
27.6157 | -35.9833 | -39.6488
27.6218 | -35.9838 | -39.654
27.6279 | -35.9842 | -39.6599
27.634 | -35.9847 | -39.6654
27.6401 | -35.9853 | -39.6711
27.6462 | -35.9857 | -39.6768
27.6523 | -35.9862 | -39.6819
27.6584 | -35.9866 | -39.6876
27.6645 | -35.9871 | -39.6929
27.6706 | -35.9877 | -39.6987
27.6767 | -35.9881 | -39.704
27.6828 | -35.9885 | -39.7092
27.6889 | -35.989 | -39.7148
27.695 | -35.9895 | -39.7205
27.7011 | -35.99 | -39.726
27.7072 | -35.9905 | -39.7321
27.7133 | -35.9909 | -39.738
27.7194 | -35.9915 | -39.744
27.7256 | -35.992 | -39.7496
27.7317 | -35.9925 | -39.7551
27.7378 | -35.9929 | -39.761
27.7439 | -35.9933 | -39.7666
27.75 | -35.9939 | -39.7725
27.7561 | -35.9943 | -39.7772
27.7622 | -35.9949 | -39.783
27.7683 | -35.9954 | -39.789
27.7744 | -35.9958 | -39.7944
27.7805 | -35.9963 | -39.7996
27.7866 | -35.9967 | -39.8058
27.7927 | -35.9971 | -39.8115
27.7988 | -35.9976 | -39.8167
27.8049 | -35.9981 | -39.8227
27.811 | -35.9985 | -39.8282
27.8171 | -35.9989 | -39.8339
27.8232 | -35.9994 | -39.8395
27.8293 | -35.9999 | -39.8455
27.8354 | -36.0003 | -39.8508
27.8415 | -36.0009 | -39.8563
27.8476 | -36.0014 | -39.8617
27.8537 | -36.0018 | -39.8671
27.8598 | -36.0022 | -39.8721
27.866 | -36.0027 | -39.8771
27.8721 | -36.0032 | -39.8826
27.8782 | -36.0037 | -39.8886
27.8843 | -36.0042 | -39.8944
27.8904 | -36.0047 | -39.9001
27.8965 | -36.0051 | -39.9056
27.9026 | -36.0055 | -39.9114
27.9087 | -36.006 | -39.9176
27.9148 | -36.0065 | -39.923
27.9209 | -36.0069 | -39.9278
27.927 | -36.0074 | -39.9332
27.9331 | -36.0078 | -39.9389
27.9392 | -36.0083 | -39.9448
27.9453 | -36.0088 | -39.9501
27.9514 | -36.0092 | -39.9556
27.9575 | -36.0097 | -39.9614
27.9636 | -36.0102 | -39.9666
27.9697 | -36.0105 | -39.972
27.9758 | -36.0109 | -39.9776
27.9819 | -36.0114 | -39.9833
27.988 | -36.0119 | -39.9886
27.9941 | -36.0124 | -39.9944
28.0002 | -36.0129 | -39.9995
28.0063 | -36.0134 | -40.0053
28.0125 | -36.0138 | -40.0111
28.0186 | -36.0143 | -40.0165
28.0247 | -36.0147 | -40.022
28.0308 | -36.015 | -40.0271
28.0369 | -36.0154 | -40.0317
28.043 | -36.0159 | -40.0369
28.0491 | -36.0164 | -40.0425
28.0552 | -36.0169 | -40.0485
28.0613 | -36.0174 | -40.0536
28.0674 | -36.0177 | -40.0591
28.0735 | -36.0182 | -40.0638

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.74 | -44.1773
21.4908 | -35.6465 | -36.5872
16.4948 | -34.9601 | -28.9004
12.6602 | -34.56 | -22.3712
9.7171 | -34.3214 | -17.2565
7.4582 | -34.1565 | -13.114
5.7244 | -34.0425 | -9.7637
4.3936 | -33.9723 | -7.0572
3.3722 | -33.9493 | -4.9279
2.5883 | -33.9407 | -3.3126
1.9866 | -33.9491 | -2.0258
1.5248 | -33.9484 | -0.99668
1.1703 | -33.9615 | -0.32175
1 | -33.9571 | 0.013412
0.89822 | -33.9628 | 0.29097
0.68941 | -33.9467 | 0.88059
0.52919 | -33.9485 | 1.3411
0.40612 | -33.9568 | 1.9502
0.31171 | -33.9461 | 2.4022
0.23928 | -33.9374 | 3.1319
0.18358 | -33.9607 | 3.544
0.14091 | -33.9625 | 4.4171
0.10819 | -33.99 | 5.5569
0.083017 | -34.0138 | 7.02
0.063705 | -34.0563 | 8.8807
0.048923 | -34.1184 | 11.3511
0.037479 | -34.2274 | 14.5162
0.028801 | -34.4071 | 18.5147
0.022078 | -34.7063 | 23.8048
0.017023 | -35.1809 | 30.1724
0.013018 | -35.9828 | 37.8312
0.010014 | -37.0806 | 45.7141

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003658 | -7.6066
0.007316 | -26.2172
0.010974 | -34.5855
0.014633 | -40.8381
0.018291 | -43.5971
0.021949 | -46.356
0.025607 | -47.5278
0.029265 | -48.284
0.032923 | -50.9878
0.036581 | -52.3024
0.04024 | -53.3225
0.043898 | -54.2745
0.047556 | -55.541
0.051214 | -56.2985
0.054872 | -57.0607
0.05853 | -57.9814
0.062188 | -58.9138
0.065847 | -59.0381
0.069505 | -60.0183
0.073163 | -60.2902
0.076821 | -60.3298
0.080479 | -61.2285
0.084137 | -62.9038
0.087796 | -62.44
0.091454 | -62.4683
0.095112 | -62.8258
0.09877 | -62.3803
0.10243 | -63.4072
0.10609 | -64.3956
0.10974 | -64.9139
0.1134 | -65.1889
0.11706 | -64.1868
0.12072 | -65.0473
0.12438 | -65.7798
0.12804 | -65.6488
0.13169 | -66.4806
0.13535 | -67.4542
0.13901 | -67.8326
0.14267 | -66.81
0.14633 | -67.8631
0.14998 | -68.0713
0.15364 | -66.4203
0.1573 | -66.6385
0.16096 | -69.161
0.16462 | -68.698
0.16828 | -68.6228
0.17193 | -68.9696
0.17559 | -68.7853
0.17925 | -69.5364
0.18291 | -69.781
0.18657 | -69.8183
0.19022 | -69.6412
0.19388 | -69.1447
0.19754 | -69.8104
0.2012 | -70.4753
0.20486 | -71.1604
0.20851 | -70.5065
0.21217 | -70.517
0.21583 | -70.7978
0.21949 | -72.3327
0.22315 | -73.2229
0.22681 | -72.1891
0.23046 | -71.2339
0.23412 | -71.7445
0.23778 | -72.9925
0.24144 | -73.5777
0.2451 | -73.5118
0.24875 | -73.5328
0.25241 | -73.2363
0.25607 | -73.0661
0.25973 | -73.1357
0.26339 | -72.8767
0.26704 | -73.5384
0.2707 | -73.618
0.27436 | -74.5322
0.27802 | -74.8918
0.28168 | -75.3683
0.28534 | -75.6276
0.28899 | -75.1442
0.29265 | -74.3351
0.29631 | -74.6014
0.29997 | -74.4958
0.30363 | -74.8188
0.30728 | -75.3322
0.31094 | -75.2555
0.3146 | -75.8334
0.31826 | -75.3695
0.32192 | -75.7046
0.32558 | -74.9036
0.32923 | -75.326
0.33289 | -76.0303
0.33655 | -76.8414
0.34021 | -76.3189
0.34387 | -75.6933
0.34752 | -76.2508
0.35118 | -77.2057
0.35484 | -76.9334
0.3585 | -76.7086
0.36216 | -77.0088
0.36582 | -77.1357
0.36947 | -76.9689
0.37313 | -76.7488
0.37679 | -77.5046
0.38045 | -78.2654
0.3841 | -77.0046
0.38776 | -76.7559
0.39142 | -77.9127
0.39508 | -78.0177
0.39874 | -78.7055
0.4024 | -78.1317
0.40605 | -78.2509
0.40971 | -77.3919
0.41337 | -76.859
0.41703 | -77.4169
0.42069 | -78.6512
0.42435 | -79.16
0.428 | -78.5973
0.43166 | -77.64
0.43532 | -78.4895
0.43898 | -79.381
0.44264 | -78.5117
0.44629 | -79.0934
0.44995 | -78.9374
0.45361 | -78.641
0.45727 | -78.4161
0.46093 | -78.0509
0.46459 | -78.5587
0.46824 | -78.6121
0.4719 | -78.6594
0.47556 | -79.4233
0.47922 | -80.4352
0.48287 | -80.1011
0.48653 | -79.7304
0.49019 | -79.4246
0.49385 | -78.8349
0.49751 | -78.3883
0.50117 | -78.9124
0.50482 | -79.2268
0.50848 | -78.821
0.51214 | -79.1492
0.5158 | -80.4541
0.51946 | -79.7645
0.52311 | -79.074
0.52677 | -78.9454
0.53043 | -78.8416
0.53409 | -79.2853
0.53775 | -80.0986
0.54141 | -80.3942
0.54506 | -80.3641
0.54872 | -80.4931
0.55238 | -80.2979
0.55604 | -79.8486
0.5597 | -79.5701
0.56336 | -79.6456
0.56701 | -79.9839
0.57067 | -80.5905
0.57433 | -81.0393
0.57799 | -80.7756
0.58164 | -80.5101
0.5853 | -80.6223
0.58896 | -80.9312
0.59262 | -81.1341
0.59628 | -81.2538
0.59994 | -81.1344
0.60359 | -80.9167
0.60725 | -80.618
0.61091 | -80.515
0.61457 | -80.9029
0.61823 | -81.2057
0.62189 | -81.0872
0.62554 | -81.0062
0.6292 | -81.0527
0.63286 | -80.9583
0.63652 | -81.0426
0.64018 | -81.5205
0.64383 | -81.8808
0.64749 | -81.8484
0.65115 | -81.3986
0.65481 | -81.0259
0.65847 | -81.0288
0.66212 | -81.1818
0.66578 | -81.4605
0.66944 | -82.2062
0.6731 | -82.9188
0.67676 | -82.9259
0.68041 | -82.4925
0.68407 | -82.1859
0.68773 | -82.0383
0.69139 | -82.0447
0.69505 | -82.1981
0.69871 | -82.7584
0.70236 | -82.8993
0.70602 | -82.5086
0.70968 | -82.3818
0.71334 | -82.5857
0.717 | -83.0534
0.72066 | -83.1016
0.72431 | -82.7405
0.72797 | -82.7228
0.73163 | -82.8281
0.73529 | -83.0064
0.73895 | -83.227
0.7426 | -83.4494
0.74626 | -83.5843
0.74992 | -83.635
0.75358 | -83.8653
0.75724 | -84.0536
0.76089 | -84.0389
0.76455 | -83.8373
0.76821 | -83.3942
0.77187 | -83.2528
0.77553 | -83.4369
0.77919 | -83.4541
0.78284 | -83.0162
0.7865 | -82.6717
0.79016 | -82.7943
0.79382 | -83.2562
0.79748 | -83.6421
0.80113 | -83.7256
0.80479 | -83.5179
0.80845 | -83.8442
0.81211 | -84.4766
0.81577 | -85.0062
0.81942 | -84.9257
0.82308 | -84.4546
0.82674 | -84.4955
0.8304 | -85.1316
0.83406 | -85.3465
0.83771 | -84.8428
0.84137 | -84.6859
0.84503 | -84.5455
0.84869 | -84.4878
0.85235 | -84.4601
0.85601 | -84.5693
0.85966 | -84.974
0.86332 | -84.9967
0.86698 | -84.7054
0.87064 | -84.7997
0.8743 | -85.0381
0.87796 | -84.9537
0.88161 | -85.0299
0.88527 | -85.5961
0.88893 | -85.6195
0.89259 | -85.1702
0.89625 | -84.7486
0.8999 | -84.6819
0.90356 | -85.103
0.90722 | -85.2348
0.91088 | -84.9244
0.91454 | -85.3714
0.91819 | -86.1262
0.92185 | -86.1971
0.92551 | -85.8846
0.92917 | -85.5644
0.93283 | -85.385
0.93649 | -85.2085
0.94014 | -85.1293
0.9438 | -85.38
0.94746 | -85.3681
0.95112 | -85.1538
0.95478 | -85.3509
0.95843 | -85.5178
0.96209 | -85.5931
0.96575 | -85.8574
0.96941 | -85.7495
0.97307 | -85.5205
0.97672 | -85.4522
0.98038 | -85.1713
0.98404 | -85.0142
0.9877 | -85.092
0.99136 | -85.6474
0.99502 | -86.0315
0.99867 | -85.6072
1.0023 | -85.4223
1.006 | -85.778
1.0096 | -85.8774
1.0133 | -86.0587
1.017 | -86.5794
1.0206 | -86.7857
1.0243 | -86.6169
1.0279 | -86.1945
1.0316 | -85.7896
1.0353 | -85.7197
1.0389 | -85.9449
1.0426 | -86.521
1.0462 | -86.7674
1.0499 | -86.4822
1.0535 | -86.195
1.0572 | -86.2561
1.0609 | -86.5724
1.0645 | -86.8106
1.0682 | -86.7468
1.0718 | -86.5023
1.0755 | -86.299
1.0792 | -86.2965
1.0828 | -86.587
1.0865 | -87.049
1.0901 | -87.2002
1.0938 | -86.7924
1.0974 | -86.3203
1.1011 | -86.2597
1.1048 | -86.5786
1.1084 | -86.9164
1.1121 | -86.8798
1.1157 | -86.5238
1.1194 | -86.2328
1.1231 | -86.2137
1.1267 | -86.4182
1.1304 | -86.7401
1.134 | -87.1589
1.1377 | -87.6427
1.1413 | -88.0014
1.145 | -87.999
1.1487 | -87.6217
1.1523 | -87.1145
1.156 | -86.7792
1.1596 | -86.8154
1.1633 | -87.1853
1.1669 | -87.5777
1.1706 | -87.7439
1.1743 | -87.752
1.1779 | -87.7981
1.1816 | -87.8848
1.1852 | -87.823
1.1889 | -87.656
1.1926 | -87.6536
1.1962 | -87.8205
1.1999 | -87.9397
1.2035 | -87.8763
1.2072 | -87.632
1.2108 | -87.4261
1.2145 | -87.4608
1.2182 | -87.6158
1.2218 | -87.693
1.2255 | -87.607
1.2291 | -87.3524
1.2328 | -87.1769
1.2365 | -87.2756
1.2401 | -87.5177
1.2438 | -87.6812
1.2474 | -87.6933
1.2511 | -87.688
1.2547 | -87.82
1.2584 | -87.9893
1.2621 | -88.0502
1.2657 | -88.0415
1.2694 | -88.0549
1.273 | -88.0428
1.2767 | -87.8339
1.2804 | -87.4153
1.284 | -87.1275
1.2877 | -87.1956
1.2913 | -87.46
1.295 | -87.6892
1.2986 | -87.9723
1.3023 | -88.3841
1.306 | -88.6707
1.3096 | -88.6382
1.3133 | -88.4989
1.3169 | -88.4382
1.3206 | -88.4446
1.3242 | -88.4238
1.3279 | -88.3112
1.3316 | -88.0732
1.3352 | -87.7391
1.3389 | -87.5361
1.3425 | -87.6879
1.3462 | -88.1072
1.3499 | -88.5142
1.3535 | -88.7213
1.3572 | -88.7641
1.3608 | -88.8403
1.3645 | -89.0268
1.3681 | -89.294
1.3718 | -89.5738
1.3755 | -89.7035
1.3791 | -89.5832
1.3828 | -89.328
1.3864 | -89.0991
1.3901 | -88.9056
1.3938 | -88.7653
1.3974 | -88.7748
1.4011 | -88.8971
1.4047 | -89.0469
1.4084 | -89.0941
1.412 | -88.9504
1.4157 | -88.8825
1.4194 | -89.1888
1.423 | -89.7469
1.4267 | -90.0748
1.4303 | -89.9321
1.434 | -89.568
1.4377 | -89.2682
1.4413 | -89.1295
1.445 | -89.0966
1.4486 | -89.1362
1.4523 | -89.3397
1.4559 | -89.7343
1.4596 | -90.1248
1.4633 | -90.2732
1.4669 | -90.1657
1.4706 | -89.8595
1.4742 | -89.4168
1.4779 | -89.0275
1.4815 | -88.8564
1.4852 | -88.9145
1.4889 | -89.1351
1.4925 | -89.4375
1.4962 | -89.7658
1.4998 | -90.0348
1.5035 | -90.0562
1.5072 | -89.7349
1.5108 | -89.2861
1.5145 | -89.0587
1.5181 | -89.1934
1.5218 | -89.607
1.5254 | -90.0914
1.5291 | -90.3559
1.5328 | -90.3091
1.5364 | -90.2022
1.5401 | -90.2019
1.5437 | -90.2113
1.5474 | -90.0948
1.5511 | -89.9167
1.5547 | -89.8475
1.5584 | -89.9142
1.562 | -89.978
1.5657 | -89.9713
1.5693 | -89.8905
1.573 | -89.7321
1.5767 | -89.6111
1.5803 | -89.8318
1.584 | -90.1042
1.5876 | -90.452
1.5913 | -90.7116
1.595 | -90.7472
1.5986 | -90.5443
1.6023 | -90.2297
1.6059 | -90.0052
1.6096 | -89.9756
1.6132 | -90.1301
1.6169 | -90.4074
1.6206 | -90.7496
1.6242 | -91.0527
1.6279 | -91.166
1.6315 | -91.0973
1.6352 | -91.0075
1.6388 | -90.9838
1.6425 | -91.0175
1.6462 | -91.0503
1.6498 | -91.0662
1.6535 | -91.0785
1.6571 | -91.1222
1.6608 | -91.2382
1.6645 | -91.404
1.6681 | -91.5516
1.6718 | -91.6446
1.6754 | -91.6844
1.6791 | -91.6551
1.6827 | -91.554
1.6864 | -91.4135
1.6901 | -91.2581
1.6937 | -91.1089
1.6974 | -91.003
1.701 | -90.9239
1.7047 | -90.8522
1.7084 | -90.8123
1.712 | -90.8199
1.7157 | -90.7894
1.7193 | -90.6839
1.723 | -90.5586
1.7266 | -90.5058
1.7303 | -90.5879
1.734 | -90.8077
1.7376 | -91.0725
1.7413 | -91.2949
1.7449 | -91.4285
1.7486 | -91.4392
1.7523 | -91.3296
1.7559 | -91.2011
1.7596 | -91.1584
1.7632 | -91.2057
1.7669 | -91.2655
1.7705 | -91.2536
1.7742 | -91.1347
1.7779 | -90.9684
1.7815 | -90.8476
1.7852 | -90.8244
1.7888 | -90.8773
1.7925 | -90.9363
1.7961 | -90.9475
1.7998 | -90.9148
1.8035 | -90.8519
1.8071 | -90.7689
1.8108 | -90.7386
1.8144 | -90.8549
1.8181 | -91.1649
1.8218 | -91.6461
1.8254 | -92.2077
1.8291 | -92.6523
1.8327 | -92.7629
1.8364 | -92.5701
1.84 | -92.2834
1.8437 | -91.9965
1.8474 | -91.7148
1.851 | -91.4467
1.8547 | -91.2577
1.8583 | -91.2438
1.862 | -91.436
1.8657 | -91.732
1.8693 | -91.9387
1.873 | -91.8794
1.8766 | -91.6137
1.8803 | -91.3443
1.8839 | -91.1816
1.8876 | -91.1202
1.8913 | -91.1392
1.8949 | -91.255
1.8986 | -91.476
1.9022 | -91.7479
1.9059 | -92.0167
1.9096 | -92.1704
1.9132 | -92.1895
1.9169 | -92.1673
1.9205 | -92.1753
1.9242 | -92.1496
1.9278 | -91.9783
1.9315 | -91.6512
1.9352 | -91.3345
1.9388 | -91.2154
1.9425 | -91.3538
1.9461 | -91.6008
1.9498 | -91.7738
1.9534 | -91.787
1.9571 | -91.6858
1.9608 | -91.5377
1.9644 | -91.412
1.9681 | -91.3498
1.9717 | -91.3769
1.9754 | -91.5075
1.9791 | -91.7089
1.9827 | -91.9195
1.9864 | -92.1151
1.99 | -92.2715
1.9937 | -92.3484
1.9973 | -92.3173
2.001 | -92.2342
2.0047 | -92.1834
2.0083 | -92.2388
2.012 | -92.3865
2.0156 | -92.519
2.0193 | -92.5365
2.023 | -92.4356
2.0266 | -92.3073
2.0303 | -92.2658
2.0339 | -92.3632
2.0376 | -92.5914
2.0412 | -92.8818
2.0449 | -93.1517
2.0486 | -93.3659
2.0522 | -93.5038
2.0559 | -93.5661
2.0595 | -93.5682
2.0632 | -93.5787
2.0669 | -93.6616
2.0705 | -93.7544
2.0742 | -93.7259
2.0778 | -93.5353
2.0815 | -93.2378
2.0851 | -92.9337
2.0888 | -92.6694
2.0925 | -92.5106
2.0961 | -92.5103
2.0998 | -92.6679
2.1034 | -92.9152
2.1071 | -93.1164
2.1107 | -93.0468
2.1144 | -92.9901
2.1181 | -92.8941
2.1217 | -92.8153
2.1254 | -92.8134
2.129 | -92.8992
2.1327 | -93.0379
2.1364 | -93.169
2.14 | -93.1999
2.1437 | -93.1327
2.1473 | -93.0506
2.151 | -93.0317
2.1546 | -93.1047
2.1583 | -93.2379
2.162 | -93.3704
2.1656 | -93.4452
2.1693 | -93.4436
2.1729 | -93.3939
2.1766 | -93.3293
2.1803 | -93.2431
2.1839 | -93.0864
2.1876 | -92.8757
2.1912 | -92.6907
2.1949 | -92.6246
2.1985 | -92.7236
2.2022 | -92.9736
2.2059 | -93.2724
2.2095 | -93.497
2.2132 | -93.5745
2.2168 | -93.5355
2.2205 | -93.4704
2.2242 | -93.3822
2.2278 | -93.249
2.2315 | -93.1136
2.2351 | -93.0134
2.2388 | -92.942
2.2424 | -92.8704
2.2461 | -92.8115
2.2498 | -92.7986
2.2534 | -92.8685
2.2571 | -93.0137
2.2607 | -93.1951
2.2644 | -93.4075
2.2681 | -93.6121
2.2717 | -93.7615
2.2754 | -93.8105
2.279 | -93.7649
2.2827 | -93.6449
2.2863 | -93.5103
2.29 | -93.4152
2.2937 | -93.3834
2.2973 | -93.4029
2.301 | -93.443
2.3046 | -93.4478
2.3083 | -93.4135
2.3119 | -93.3374
2.3156 | -93.2267
2.3193 | -93.0899
2.3229 | -92.9812
2.3266 | -92.9504
2.3302 | -93.0108
2.3339 | -93.1176
2.3376 | -93.2016
2.3412 | -93.2227
2.3449 | -93.2065
2.3485 | -93.1796
2.3522 | -93.1799
2.3558 | -93.2095
2.3595 | -93.2369
2.3632 | -93.222
2.3668 | -93.1503
2.3705 | -93.0477
2.3741 | -92.9685
2.3778 | -92.9663
2.3815 | -93.0567
2.3851 | -93.1964
2.3888 | -93.3075
2.3924 | -93.314
2.3961 | -93.2321
2.3997 | -93.1491
2.4034 | -93.1429
2.4071 | -93.2297
2.4107 | -93.3624
2.4144 | -93.4675
2.418 | -93.5144
2.4217 | -93.5088
2.4254 | -93.4899
2.429 | -93.491
2.4327 | -93.5243
2.4363 | -93.5802
2.44 | -93.6396
2.4436 | -93.7029
2.4473 | -93.7857
2.451 | -93.8733
2.4546 | -93.9448
2.4583 | -93.9714
2.4619 | -93.9713
2.4656 | -93.9755
2.4692 | -93.9759
2.4729 | -93.9466
2.4766 | -93.8867
2.4802 | -93.8333
2.4839 | -93.8222
2.4875 | -93.8469
2.4912 | -93.8745
2.4949 | -93.885
2.4985 | -93.8774
2.5022 | -93.8496
2.5058 | -93.8287
2.5095 | -93.8679
2.5131 | -93.9807
2.5168 | -94.1103
2.5205 | -94.198
2.5241 | -94.2188
2.5278 | -94.212
2.5314 | -94.2353
2.5351 | -94.2972
2.5388 | -94.384
2.5424 | -94.4394
2.5461 | -94.4142
2.5497 | -94.2925
2.5534 | -94.1121
2.557 | -93.9001
2.5607 | -93.6873
2.5644 | -93.5081
2.568 | -93.3995
2.5717 | -93.4015
2.5753 | -93.545
2.579 | -93.8195
2.5827 | -94.1406
2.5863 | -94.3906
2.59 | -94.5054
2.5936 | -94.4819
2.5973 | -94.3796
2.6009 | -94.27
2.6046 | -94.2433
2.6083 | -94.3449
2.6119 | -94.5597
2.6156 | -94.7937
2.6192 | -94.9129
2.6229 | -94.8694
2.6265 | -94.7283
2.6302 | -94.6051
2.6339 | -94.5633
2.6375 | -94.6261
2.6412 | -94.6789
2.6448 | -94.7594
2.6485 | -94.862
2.6522 | -94.9752
2.6558 | -95.0739
2.6595 | -95.1326
2.6631 | -95.1456
2.6668 | -95.1339
2.6704 | -95.1027
2.6741 | -95.07
2.6778 | -95.0403
2.6814 | -95.0151
2.6851 | -94.9665
2.6887 | -94.8915
2.6924 | -94.8053
2.6961 | -94.71
2.6997 | -94.648
2.7034 | -94.6551
2.707 | -94.7485
2.7107 | -94.9231
2.7143 | -95.1344
2.718 | -95.3404
2.7217 | -95.5019
2.7253 | -95.6141
2.729 | -95.6687
2.7326 | -95.6648
2.7363 | -95.6148
2.74 | -95.5357
2.7436 | -95.4479
2.7473 | -95.3404
2.7509 | -95.211
2.7546 | -95.0453
2.7582 | -94.8389
2.7619 | -94.62
2.7656 | -94.4402
2.7692 | -94.3343
2.7729 | -94.3295
2.7765 | -94.4252
2.7802 | -94.592
2.7838 | -94.7801
2.7875 | -94.9453
2.7912 | -95.0674
2.7948 | -95.1355
2.7985 | -95.1593
2.8021 | -95.1458
2.8058 | -95.1217
2.8095 | -95.1086
2.8131 | -95.1128
2.8168 | -95.1288
2.8204 | -95.1224
2.8241 | -95.0851
2.8277 | -95.0207
2.8314 | -94.9632
2.8351 | -94.9326
2.8387 | -94.9333
2.8424 | -94.9549
2.846 | -94.9743
2.8497 | -94.994
2.8534 | -94.9978
2.857 | -94.9759
2.8607 | -94.9154
2.8643 | -94.8407
2.868 | -94.7912
2.8716 | -94.7964
2.8753 | -94.8746
2.879 | -95.0221
2.8826 | -95.2178
2.8863 | -95.4261
2.8899 | -95.6049
2.8936 | -95.7229
2.8973 | -95.7746
2.9009 | -95.7896
2.9046 | -95.7863
2.9082 | -95.7715
2.9119 | -95.7307
2.9155 | -95.6846
2.9192 | -95.6483
2.9229 | -95.6337
2.9265 | -95.6425
2.9302 | -95.6757
2.9338 | -95.7335
2.9375 | -95.8332
2.9411 | -95.9488
2.9448 | -96.0367
2.9485 | -96.0574
2.9521 | -96.0113
2.9558 | -95.9162
2.9594 | -95.8049
2.9631 | -95.6943
2.9668 | -95.6095
2.9704 | -95.5556
2.9741 | -95.5294
2.9777 | -95.5419
2.9814 | -95.5948
2.985 | -95.6756
2.9887 | -95.7838
2.9924 | -95.896
2.996 | -95.9887
2.9997 | -96.0329
3.0033 | -96.013
3.007 | -95.9288
3.0107 | -95.7756
3.0143 | -95.5761
3.018 | -95.3735
3.0216 | -95.2112
3.0253 | -95.1097
3.0289 | -95.0657
3.0326 | -95.0611
3.0363 | -95.0672
3.0399 | -95.0527
3.0436 | -95.0177
3.0472 | -94.9748
3.0509 | -94.9407
3.0546 | -94.9276
3.0582 | -94.9611
3.0619 | -95.0439
3.0655 | -95.1761
3.0692 | -95.3571
3.0728 | -95.5572
3.0765 | -95.7446
3.0802 | -95.8961
3.0838 | -95.9921
3.0875 | -96.0413
3.0911 | -96.0275
3.0948 | -95.958
3.0985 | -95.8522
3.1021 | -95.7368
3.1058 | -95.6364
3.1094 | -95.582
3.1131 | -95.5918
3.1167 | -95.6578
3.1204 | -95.7478
3.1241 | -95.8335
3.1277 | -95.8886
3.1314 | -95.9055
3.135 | -95.8983
3.1387 | -95.8862
3.1423 | -95.877
3.146 | -95.8629
3.1497 | -95.8491
3.1533 | -95.8549
3.157 | -95.8922
3.1606 | -95.9695
3.1643 | -96.0742
3.168 | -96.1568
3.1716 | -96.2236
3.1753 | -96.251
3.1789 | -96.239
3.1826 | -96.1814
3.1862 | -96.083
3.1899 | -95.9682
3.1936 | -95.8655
3.1972 | -95.8086
3.2009 | -95.8178
3.2045 | -95.8846
3.2082 | -95.9879
3.2119 | -96.0928
3.2155 | -96.1863
3.2192 | -96.2481
3.2228 | -96.2852
3.2265 | -96.3218
3.2301 | -96.365
3.2338 | -96.406
3.2375 | -96.444
3.2411 | -96.4583
3.2448 | -96.4447
3.2484 | -96.4187
3.2521 | -96.3967
3.2557 | -96.3954
3.2594 | -96.4253
3.2631 | -96.4777
3.2667 | -96.5356
3.2704 | -96.5791
3.274 | -96.6016
3.2777 | -96.604
3.2814 | -96.5892
3.285 | -96.5499
3.2887 | -96.495
3.2923 | -96.4219
3.296 | -96.3413
3.2996 | -96.2722
3.3033 | -96.2329
3.307 | -96.2348
3.3106 | -96.2743
3.3143 | -96.3361
3.3179 | -96.3979
3.3216 | -96.4442
3.3253 | -96.4525
3.3289 | -96.4203
3.3326 | -96.3708
3.3362 | -96.3231
3.3399 | -96.3055
3.3435 | -96.3141
3.3472 | -96.3516
3.3509 | -96.4213
3.3545 | -96.5016
3.3582 | -96.5797
3.3618 | -96.6502
3.3655 | -96.7098
3.3692 | -96.7558
3.3728 | -96.7804
3.3765 | -96.7822
3.3801 | -96.7629
3.3838 | -96.7168
3.3874 | -96.6279
3.3911 | -96.5134
3.3948 | -96.3868
3.3984 | -96.2727
3.4021 | -96.1873
3.4057 | -96.1419
3.4094 | -96.1458
3.4131 | -96.1936
3.4167 | -96.2699
3.4204 | -96.3569
3.424 | -96.4387
3.4277 | -96.4939
3.4313 | -96.5051
3.435 | -96.4925
3.4387 | -96.4928
3.4423 | -96.5353
3.446 | -96.6325
3.4496 | -96.7698
3.4533 | -96.9051
3.4569 | -96.9877
3.4606 | -96.9836
3.4643 | -96.9011
3.4679 | -96.766
3.4716 | -96.6112
3.4752 | -96.4637
3.4789 | -96.3493
3.4826 | -96.2801
3.4862 | -96.2544
3.4899 | -96.2724
3.4935 | -96.3219
3.4972 | -96.3881
3.5008 | -96.4581
3.5045 | -96.5311
3.5082 | -96.5932
3.5118 | -96.6472
3.5155 | -96.7003
3.5191 | -96.7438
3.5228 | -96.7669
3.5265 | -96.7868
3.5301 | -96.8121
3.5338 | -96.8362
3.5374 | -96.8602
3.5411 | -96.8769
3.5447 | -96.874
3.5484 | -96.8437
3.5521 | -96.8018
3.5557 | -96.7548
3.5594 | -96.713
3.563 | -96.6903
3.5667 | -96.7013
3.5704 | -96.7561
3.574 | -96.8623
3.5777 | -97.0064
3.5813 | -97.1427
3.585 | -97.2383
3.5886 | -97.2568
3.5923 | -97.1843
3.596 | -97.0588
3.5996 | -96.9156
3.6033 | -96.7977
3.6069 | -96.7218
3.6106 | -96.691
3.6142 | -96.6945
3.6179 | -96.7196
3.6216 | -96.7489
3.6252 | -96.7668
3.6289 | -96.7587
3.6325 | -96.7289
3.6362 | -96.6767
3.6399 | -96.6099
3.6435 | -96.5552
3.6472 | -96.5194
3.6508 | -96.5039
3.6545 | -96.5068
3.6581 | -96.519
3.6618 | -96.531
3.6655 | -96.5359
3.6691 | -96.5332
3.6728 | -96.521
3.6764 | -96.5015
3.6801 | -96.4987
3.6838 | -96.5272
3.6874 | -96.5921
3.6911 | -96.7551
3.6947 | -96.8883
3.6984 | -97.0418
3.702 | -97.1931
3.7057 | -97.3155
3.7094 | -97.3903
3.713 | -97.4121
3.7167 | -97.4022
3.7203 | -97.361
3.724 | -97.3078
3.7277 | -97.2519
3.7313 | -97.2034
3.735 | -97.1735
3.7386 | -97.1661
3.7423 | -97.1876
3.7459 | -97.2403
3.7496 | -97.314
3.7533 | -97.397
3.7569 | -97.489
3.7606 | -97.5817
3.7642 | -97.6645
3.7679 | -97.7396
3.7715 | -97.799
3.7752 | -97.8452
3.7789 | -97.848
3.7825 | -97.8267
3.7862 | -97.7878
3.7898 | -97.7407
3.7935 | -97.6929
3.7972 | -97.6558
3.8008 | -97.6365
3.8045 | -97.6347
3.8081 | -97.6399
3.8118 | -97.6519
3.8154 | -97.6684
3.8191 | -97.6789
3.8228 | -97.6862
3.8264 | -97.6973
3.8301 | -97.7048
3.8337 | -97.7106
3.8374 | -97.7209
3.8411 | -97.742
3.8447 | -97.7601
3.8484 | -97.7713
3.852 | -97.7695
3.8557 | -97.7457
3.8593 | -97.7142
3.863 | -97.6658
3.8667 | -97.6137
3.8703 | -97.5614
3.874 | -97.5149
3.8776 | -97.4838
3.8813 | -97.4781
3.885 | -97.5096
3.8886 | -97.5764
3.8923 | -97.676
3.8959 | -97.7935
3.8996 | -97.9042
3.9032 | -97.99
3.9069 | -98.0344
3.9106 | -98.0334
3.9142 | -98.0075
3.9179 | -97.9702
3.9215 | -97.924
3.9252 | -97.903
3.9288 | -97.8994
3.9325 | -97.9163
3.9362 | -97.9571
3.9398 | -98.0001
3.9435 | -98.0454
3.9471 | -98.102
3.9508 | -98.1652
3.9545 | -98.224
3.9581 | -98.2738
3.9618 | -98.3174
3.9654 | -98.3478
3.9691 | -98.353
3.9727 | -98.3349
3.9764 | -98.2824
3.9801 | -98.206
3.9837 | -98.109
3.9874 | -98.0004
3.991 | -97.8941
3.9947 | -97.7986
3.9984 | -97.724
4.002 | -97.6835
4.0057 | -97.6513
4.0093 | -97.6143
4.013 | -97.6021
4.0166 | -97.6215
4.0203 | -97.668
4.024 | -97.7342
4.0276 | -97.8174
4.0313 | -97.9053
4.0349 | -97.966
4.0386 | -97.991
4.0423 | -97.9721
4.0459 | -97.9306
4.0496 | -97.8861
4.0532 | -97.8652
4.0569 | -97.8894
4.0605 | -97.9428
4.0642 | -98.0282
4.0679 | -98.1366
4.0715 | -98.2371
4.0752 | -98.3017
4.0788 | -98.3267
4.0825 | -98.3172
4.0861 | -98.2812
4.0898 | -98.2147
4.0935 | -98.1291
4.0971 | -98.0583
4.1008 | -98.0142
4.1044 | -97.9976
4.1081 | -98.0075
4.1118 | -98.0349
4.1154 | -98.0578
4.1191 | -98.0589
4.1227 | -98.0589
4.1264 | -98.0513
4.13 | -98.0478
4.1337 | -98.0594
4.1374 | -98.0811
4.141 | -98.1178
4.1447 | -98.1686
4.1483 | -98.2304
4.152 | -98.3007
4.1557 | -98.3779
4.1593 | -98.4274
4.163 | -98.4535
4.1666 | -98.4521
4.1703 | -98.4012
4.1739 | -98.311
4.1776 | -98.193
4.1813 | -98.0741
4.1849 | -97.963
4.1886 | -97.8846
4.1922 | -97.8458
4.1959 | -97.8603
4.1996 | -97.9251
4.2032 | -98.0279
4.2069 | -98.1499
4.2105 | -98.2679
4.2142 | -98.3682
4.2178 | -98.4243
4.2215 | -98.3584
4.2252 | -98.3286
4.2288 | -98.2529
4.2325 | -98.1506
4.2361 | -98.0401
4.2398 | -97.9274
4.2435 | -97.8238
4.2471 | -97.7339
4.2508 | -97.6673
4.2544 | -97.6337
4.2581 | -97.6285
4.2617 | -97.6463
4.2654 | -97.6866
4.2691 | -97.7387
4.2727 | -97.8007
4.2764 | -97.865
4.28 | -97.9235
4.2837 | -97.9808
4.2873 | -98.0299
4.291 | -98.0674
4.2947 | -98.1079
4.2983 | -98.153
4.302 | -98.189
4.3056 | -98.2196
4.3093 | -98.2468
4.313 | -98.2599
4.3166 | -98.2666
4.3203 | -98.2854
4.3239 | -98.3219
4.3276 | -98.3874
4.3312 | -98.475
4.3349 | -98.5797
4.3386 | -98.6991
4.3422 | -98.8021
4.3459 | -98.8811
4.3495 | -98.9316
4.3532 | -98.9447
4.3569 | -98.9405
4.3605 | -98.9173
4.3642 | -98.8781
4.3678 | -98.8292
4.3715 | -98.783
4.3751 | -98.7386
4.3788 | -98.6973
4.3825 | -98.669
4.3861 | -98.6414
4.3898 | -98.6221
4.3934 | -98.6143
4.3971 | -98.6246
4.4008 | -98.6592
4.4044 | -98.7157
4.4081 | -98.784
4.4117 | -98.8796
4.4154 | -98.995
4.419 | -99.1193
4.4227 | -99.2477
4.4264 | -99.3605
4.43 | -99.4486
4.4337 | -99.493
4.4373 | -99.5035
4.441 | -99.4814
4.4446 | -99.4385
4.4483 | -99.3777
4.452 | -99.3032
4.4556 | -99.2259
4.4593 | -99.145
4.4629 | -99.0538
4.4666 | -98.9659
4.4703 | -98.8813
4.4739 | -98.808
4.4776 | -98.7496
4.4812 | -98.7011
4.4849 | -98.6643
4.4885 | -98.6518
4.4922 | -98.6619
4.4959 | -98.6931
4.4995 | -98.7472
4.5032 | -98.8203
4.5068 | -98.9004
4.5105 | -98.9749
4.5142 | -99.0359
4.5178 | -99.0817
4.5215 | -99.1091
4.5251 | -99.1075
4.5288 | -99.0811
4.5324 | -99.0465
4.5361 | -98.9961
4.5398 | -98.9442
4.5434 | -98.8974
4.5471 | -98.8656
4.5507 | -98.8453
4.5544 | -98.8387
4.558 | -98.8307
4.5617 | -98.8092
4.5654 | -98.7846
4.569 | -98.7492
4.5727 | -98.7045
4.5763 | -98.6483
4.58 | -98.5848
4.5837 | -98.5138
4.5873 | -98.4493
4.591 | -98.4003
4.5946 | -98.3796
4.5983 | -98.3893
4.6019 | -98.435
4.6056 | -98.5151
4.6093 | -98.6171
4.6129 | -98.7298
4.6166 | -98.841
4.6202 | -98.9333
4.6239 | -98.9916
4.6276 | -99.0225
4.6312 | -99.0262
4.6349 | -98.9938
4.6385 | -98.9437
4.6422 | -98.8874
4.6458 | -98.8445
4.6495 | -98.8253
4.6532 | -98.8301
4.6568 | -98.8657
4.6605 | -98.9314
4.6641 | -99.021
4.6678 | -99.1251
4.6715 | -99.2284
4.6751 | -99.3218
4.6788 | -99.3944
4.6824 | -99.4373
4.6861 | -99.4488
4.6897 | -99.438
4.6934 | -99.4017
4.6971 | -99.3416
4.7007 | -99.2849
4.7044 | -99.2313
4.708 | -99.1832
4.7117 | -99.1618
4.7153 | -99.1624
4.719 | -99.1813
4.7227 | -99.2187
4.7263 | -99.2575
4.73 | -99.2997
4.7336 | -99.3419
4.7373 | -99.3811
4.741 | -99.4129
4.7446 | -99.4348
4.7483 | -99.4401
4.7519 | -99.4551
4.7556 | -99.4664
4.7592 | -99.4838
4.7629 | -99.5134
4.7666 | -99.55
4.7702 | -99.5946
4.7739 | -99.6342
4.7775 | -99.6716
4.7812 | -99.6947
4.7849 | -99.688
4.7885 | -99.6605
4.7922 | -99.6247
4.7958 | -99.5797
4.7995 | -99.5272
4.8031 | -99.4677
4.8068 | -99.4114
4.8105 | -99.3633
4.8141 | -99.3277
4.8178 | -99.3035
4.8214 | -99.2956
4.8251 | -99.3053
4.8288 | -99.3248
4.8324 | -99.348
4.8361 | -99.3732
4.8397 | -99.4038
4.8434 | -99.4351
4.847 | -99.4637
4.8507 | -99.486
4.8544 | -99.5012
4.858 | -99.5139
4.8617 | -99.5187
4.8653 | -99.5232
4.869 | -99.521
4.8727 | -99.5226
4.8763 | -99.5338
4.88 | -99.5518
4.8836 | -99.5705
4.8873 | -99.5948
4.8909 | -99.626
4.8946 | -99.6629
4.8983 | -99.7024
4.9019 | -99.7447
4.9056 | -99.7873
4.9092 | -99.8301
4.9129 | -99.8595
4.9165 | -99.8799
4.9202 | -99.9
4.9239 | -99.9093
4.9275 | -99.9062
4.9312 | -99.8972
4.9348 | -99.8819
4.9385 | -99.8716
4.9422 | -99.8622
4.9458 | -99.8636
4.9495 | -99.8682
4.9531 | -99.8744
4.9568 | -99.8813
4.9604 | -99.8802
4.9641 | -99.8635
4.9678 | -99.8352
4.9714 | -99.7883
4.9751 | -99.7328
4.9787 | -99.6748
4.9824 | -99.617
4.9861 | -99.5597
4.9897 | -99.5039
4.9934 | -99.4539
4.997 | -99.4108
5.0007 | -99.3802
5.0043 | -99.3658
5.008 | -99.3545
5.0117 | -99.3602
5.0153 | -99.3782
5.019 | -99.4021
5.0226 | -99.4266
5.0263 | -99.4542
5.03 | -99.4771
5.0336 | -99.4911
5.0373 | -99.5089
5.0409 | -99.5328
5.0446 | -99.5658
5.0482 | -99.6197
5.0519 | -99.6822
5.0556 | -99.7504
5.0592 | -99.8235
5.0629 | -99.9015
5.0665 | -99.9722
5.0702 | -100.0363
5.0738 | -100.0989
5.0775 | -100.1579
5.0812 | -100.2143
5.0848 | -100.2811
5.0885 | -100.3381
5.0921 | -100.3797
5.0958 | -100.4126
5.0995 | -100.4326
5.1031 | -100.4289
5.1068 | -100.4162
5.1104 | -100.3996
5.1141 | -100.368
5.1177 | -100.3364
5.1214 | -100.3136
5.1251 | -100.2956
5.1287 | -100.2804
5.1324 | -100.2719
5.136 | -100.262
5.1397 | -100.2499
5.1434 | -100.2368
5.147 | -100.2121
5.1507 | -100.1828
5.1543 | -100.1539
5.158 | -100.1297
5.1616 | -100.1044
5.1653 | -100.0875
5.169 | -100.0853
5.1726 | -100.0919
5.1763 | -100.1028
5.1799 | -100.1263
5.1836 | -100.1467
5.1873 | -100.1644
5.1909 | -100.1813
5.1946 | -100.1901
5.1982 | -100.1881
5.2019 | -100.1814
5.2055 | -100.177
5.2092 | -100.1624
5.2129 | -100.1433
5.2165 | -100.1253
5.2202 | -100.1162
5.2238 | -100.1227
5.2275 | -100.1376
5.2311 | -100.1657
5.2348 | -100.1972
5.2385 | -100.2298
5.2421 | -100.2729
5.2458 | -100.3148
5.2494 | -100.3525
5.2531 | -100.3789
5.2568 | -100.3891
5.2604 | -100.3907
5.2641 | -100.3799
5.2677 | -100.3551
5.2714 | -100.3223
5.275 | -100.3047
5.2787 | -100.2895
5.2824 | -100.2773
5.286 | -100.274
5.2897 | -100.284
5.2933 | -100.2997
5.297 | -100.3094
5.3007 | -100.3104
5.3043 | -100.3015
5.308 | -100.2877
5.3116 | -100.2661
5.3153 | -100.244
5.3189 | -100.2196
5.3226 | -100.1881
5.3263 | -100.1609
5.3299 | -100.1389
5.3336 | -100.1306
5.3372 | -100.1305
5.3409 | -100.1281
5.3446 | -100.1323
5.3482 | -100.1464
5.3519 | -100.1653
5.3555 | -100.1827
5.3592 | -100.2014
5.3628 | -100.2166
5.3665 | -100.2292
5.3702 | -100.2339
5.3738 | -100.2383
5.3775 | -100.2446
5.3811 | -100.251
5.3848 | -100.2578
5.3884 | -100.2622
5.3921 | -100.2672
5.3958 | -100.2703
5.3994 | -100.2709
5.4031 | -100.2699
5.4067 | -100.2627
5.4104 | -100.2551
5.4141 | -100.248
5.4177 | -100.2448
5.4214 | -100.2483
5.425 | -100.2729
5.4287 | -100.3103
5.4323 | -100.3581
5.436 | -100.4175
5.4397 | -100.4845
5.4433 | -100.5513
5.447 | -100.6117
5.4506 | -100.6684
5.4543 | -100.7165
5.458 | -100.7565
5.4616 | -100.7891
5.4653 | -100.8056
5.4689 | -100.8126
5.4726 | -100.809
5.4762 | -100.787
5.4799 | -100.7418
5.4836 | -100.6887
5.4872 | -100.6215
5.4909 | -100.5494
5.4945 | -100.4725
5.4982 | -100.3999
5.5019 | -100.328
5.5055 | -100.2729
5.5092 | -100.2346
5.5128 | -100.2131
5.5165 | -100.2048
5.5201 | -100.2186
5.5238 | -100.2467
5.5275 | -100.2833
5.5311 | -100.3298
5.5348 | -100.3776
5.5384 | -100.4237
5.5421 | -100.4617
5.5457 | -100.4935
5.5494 | -100.5261
5.5531 | -100.5528
5.5567 | -100.5728
5.5604 | -100.59
5.564 | -100.6144
5.5677 | -100.6296
5.5714 | -100.6523
5.575 | -100.6817
5.5787 | -100.7218
5.5823 | -100.7673
5.586 | -100.8176
5.5896 | -100.8643
5.5933 | -100.9002
5.597 | -100.9297
5.6006 | -100.9561
5.6043 | -100.9632
5.6079 | -100.9607
5.6116 | -100.9447
5.6153 | -100.9164
5.6189 | -100.8769
5.6226 | -100.8304
5.6262 | -100.7838
5.6299 | -100.7439
5.6335 | -100.7132
5.6372 | -100.6937
5.6409 | -100.691
5.6445 | -100.7004
5.6482 | -100.7151
5.6518 | -100.7444
5.6555 | -100.7837
5.6592 | -100.8213
5.6628 | -100.8454
5.6665 | -100.8638
5.6701 | -100.8721
5.6738 | -100.8703
5.6774 | -100.8569
5.6811 | -100.8302
5.6848 | -100.7916
5.6884 | -100.7534
5.6921 | -100.7141
5.6957 | -100.6777
5.6994 | -100.6505
5.7031 | -100.6369
5.7067 | -100.6357
5.7104 | -100.6469
5.714 | -100.6669
5.7177 | -100.6922
5.7213 | -100.7257
5.725 | -100.7634
5.7287 | -100.7992
5.7323 | -100.8241
5.736 | -100.8258
5.7396 | -100.8061
5.7433 | -100.7864
5.7469 | -100.7606
5.7506 | -100.7187
5.7543 | -100.6776
5.7579 | -100.6339
5.7616 | -100.6017
5.7652 | -100.5776
5.7689 | -100.5684
5.7726 | -100.57
5.7762 | -100.5885
5.7799 | -100.6217
5.7835 | -100.6586
5.7872 | -100.6977
5.7908 | -100.7437
5.7945 | -100.7872
5.7982 | -100.821
5.8018 | -100.8264
5.8055 | -100.8436
5.8091 | -100.8586
5.8128 | -100.8657
5.8165 | -100.8738
5.8201 | -100.8837
5.8238 | -100.8958
5.8274 | -100.9155
5.8311 | -100.9315
5.8347 | -100.9543
5.8384 | -100.9839
5.8421 | -101.0131
5.8457 | -101.0448
5.8494 | -101.0736
5.853 | -101.1001
5.8567 | -101.1292
5.8604 | -101.1601
5.864 | -101.1895
5.8677 | -101.2215
5.8713 | -101.2586
5.875 | -101.2952
5.8786 | -101.3237
5.8823 | -101.3543
5.886 | -101.3921
5.8896 | -101.4311
5.8933 | -101.4682
5.8969 | -101.5052
5.9006 | -101.5341
5.9042 | -101.5641
5.9079 | -101.5935
5.9116 | -101.6185
5.9152 | -101.6366
5.9189 | -101.6455
5.9225 | -101.652
5.9262 | -101.6579
5.9299 | -101.6635
5.9335 | -101.669
5.9372 | -101.6757
5.9408 | -101.6808
5.9445 | -101.6848
5.9481 | -101.6785
5.9518 | -101.6752
5.9555 | -101.6707
5.9591 | -101.6545
5.9628 | -101.6305
5.9664 | -101.6099
5.9701 | -101.5781
5.9738 | -101.5406
5.9774 | -101.498
5.9811 | -101.4551
5.9847 | -101.4102
5.9884 | -101.3687
5.992 | -101.335
5.9957 | -101.3072
5.9994 | -101.2794
6.003 | -101.261
6.0067 | -101.2451
6.0103 | -101.2314
6.014 | -101.2205
6.0177 | -101.2173
6.0213 | -101.2178
6.025 | -101.2198
6.0286 | -101.2282
6.0323 | -101.2415
6.0359 | -101.2539
6.0396 | -101.2678
6.0433 | -101.2747
6.0469 | -101.2776
6.0506 | -101.2782
6.0542 | -101.2745
6.0579 | -101.2634
6.0615 | -101.2568
6.0652 | -101.2462
6.0689 | -101.2343
6.0725 | -101.2279
6.0762 | -101.228
6.0798 | -101.2239
6.0835 | -101.2213
6.0872 | -101.2208
6.0908 | -101.2194
6.0945 | -101.2124
6.0981 | -101.2075
6.1018 | -101.2021
6.1054 | -101.1953
6.1091 | -101.1918
6.1128 | -101.1976
6.1164 | -101.2097
6.1201 | -101.2256
6.1237 | -101.2444
6.1274 | -101.2689
6.1311 | -101.2884
6.1347 | -101.3053
6.1384 | -101.3203
6.142 | -101.3321
6.1457 | -101.3332
6.1493 | -101.3237
6.153 | -101.3057
6.1567 | -101.2816
6.1603 | -101.2619
6.164 | -101.2418
6.1676 | -101.2188
6.1713 | -101.205
6.175 | -101.1982
6.1786 | -101.1964
6.1823 | -101.2085
6.1859 | -101.2287
6.1896 | -101.2486
6.1932 | -101.2723
6.1969 | -101.3032
6.2006 | -101.3314
6.2042 | -101.3599
6.2079 | -101.3989
6.2115 | -101.4384
6.2152 | -101.4801
6.2188 | -101.5164
6.2225 | -101.5561
6.2262 | -101.5941
6.2298 | -101.6323
6.2335 | -101.6732
6.2371 | -101.717
6.2408 | -101.7591
6.2445 | -101.7981
6.2481 | -101.8323
6.2518 | -101.8552
6.2554 | -101.8714
6.2591 | -101.882
6.2627 | -101.8756
6.2664 | -101.8514
6.2701 | -101.824
6.2737 | -101.7927
6.2774 | -101.7637
6.281 | -101.7448
6.2847 | -101.725
6.2884 | -101.7042
6.292 | -101.7002
6.2957 | -101.7035
6.2993 | -101.7039
6.303 | -101.7123
6.3066 | -101.7247
6.3103 | -101.7397
6.314 | -101.7497
6.3176 | -101.7552
6.3213 | -101.7516
6.3249 | -101.7483
6.3286 | -101.7428
6.3323 | -101.7278
6.3359 | -101.7224
6.3396 | -101.7129
6.3432 | -101.7044
6.3469 | -101.6974
6.3505 | -101.6977
6.3542 | -101.7049
6.3579 | -101.7139
6.3615 | -101.7323
6.3652 | -101.7503
6.3688 | -101.7681
6.3725 | -101.7821
6.3761 | -101.7901
6.3798 | -101.8052
6.3835 | -101.8211
6.3871 | -101.8281
6.3908 | -101.8342
6.3944 | -101.8392
6.3981 | -101.8358
6.4018 | -101.8244
6.4054 | -101.8106
6.4091 | -101.7961
6.4127 | -101.7798
6.4164 | -101.76
6.42 | -101.7403
6.4237 | -101.7243
6.4274 | -101.7199
6.431 | -101.719
6.4347 | -101.7176
6.4383 | -101.7225
6.442 | -101.7314
6.4457 | -101.7473
6.4493 | -101.7614
6.453 | -101.777
6.4566 | -101.7912
6.4603 | -101.7962
6.4639 | -101.7943
6.4676 | -101.7842
6.4713 | -101.7707
6.4749 | -101.7539
6.4786 | -101.7277
6.4822 | -101.704
6.4859 | -101.6857
6.4896 | -101.6744
6.4932 | -101.6722
6.4969 | -101.6824
6.5005 | -101.7027
6.5042 | -101.7309
6.5078 | -101.76
6.5115 | -101.7902
6.5152 | -101.8212
6.5188 | -101.847
6.5225 | -101.8714
6.5261 | -101.8882
6.5298 | -101.9017
6.5334 | -101.9135
6.5371 | -101.9133
6.5408 | -101.9012
6.5444 | -101.8824
6.5481 | -101.8647
6.5517 | -101.8423
6.5554 | -101.8244
6.5591 | -101.8134
6.5627 | -101.8057
6.5664 | -101.7962
6.57 | -101.7904
6.5737 | -101.7879
6.5773 | -101.7861
6.581 | -101.786
6.5847 | -101.7882
6.5883 | -101.7856
6.592 | -101.7758
6.5956 | -101.7604
6.5993 | -101.7462
6.603 | -101.7297
6.6066 | -101.7089
6.6103 | -101.6963
6.6139 | -101.6832
6.6176 | -101.6685
6.6212 | -101.6582
6.6249 | -101.6484
6.6286 | -101.6414
6.6322 | -101.6407
6.6359 | -101.6478
6.6395 | -101.6566
6.6432 | -101.6665
6.6469 | -101.6849
6.6505 | -101.7125
6.6542 | -101.7419
6.6578 | -101.7806
6.6615 | -101.8247
6.6651 | -101.8825
6.6688 | -101.9509
6.6725 | -102.0183
6.6761 | -102.0892
6.6798 | -102.162
6.6834 | -102.2339
6.6871 | -102.3059
6.6907 | -102.3787
6.6944 | -102.434
6.6981 | -102.4672
6.7017 | -102.4954
6.7054 | -102.5126
6.709 | -102.4989
6.7127 | -102.4726
6.7164 | -102.4288
6.72 | -102.3693
6.7237 | -102.2958
6.7273 | -102.2147
6.731 | -102.123
6.7346 | -102.0304
6.7383 | -101.9398
6.742 | -101.8521
6.7456 | -101.7642
6.7493 | -101.6823
6.7529 | -101.6138
6.7566 | -101.5533
6.7603 | -101.5169
6.7639 | -101.492
6.7676 | -101.4761
6.7712 | -101.4778
6.7749 | -101.4944
6.7785 | -101.5203
6.7822 | -101.5582
6.7859 | -101.6047
6.7895 | -101.6541
6.7932 | -101.7067
6.7968 | -101.7579
6.8005 | -101.8029
6.8042 | -101.8536
6.8078 | -101.9023
6.8115 | -101.9454
6.8151 | -101.9908
6.8188 | -102.0306
6.8224 | -102.0706
6.8261 | -102.1044
6.8298 | -102.1372
6.8334 | -102.1657
6.8371 | -102.1964
6.8407 | -102.2329
6.8444 | -102.2704
6.848 | -102.3075
6.8517 | -102.3395
6.8554 | -102.3621
6.859 | -102.3703
6.8627 | -102.3804
6.8663 | -102.3857
6.87 | -102.3898
6.8737 | -102.3916
6.8773 | -102.3802
6.881 | -102.3661
6.8846 | -102.3529
6.8883 | -102.3448
6.8919 | -102.3415
6.8956 | -102.343
6.8993 | -102.3502
6.9029 | -102.3615
6.9066 | -102.3707
6.9102 | -102.3859
6.9139 | -102.4075
6.9176 | -102.4279
6.9212 | -102.4495
6.9249 | -102.4683
6.9285 | -102.485
6.9322 | -102.4949
6.9358 | -102.4976
6.9395 | -102.493
6.9432 | -102.4818
6.9468 | -102.4677
6.9505 | -102.452
6.9541 | -102.43
6.9578 | -102.4037
6.9615 | -102.3704
6.9651 | -102.3359
6.9688 | -102.3048
6.9724 | -102.2763
6.9761 | -102.2484
6.9797 | -102.2298
6.9834 | -102.2249
6.9871 | -102.2149
6.9907 | -102.2071
6.9944 | -102.2022
6.998 | -102.2
7.0017 | -102.2001
7.0053 | -102.1995
7.009 | -102.1967
7.0127 | -102.2001
7.0163 | -102.2055
7.02 | -102.21
7.0236 | -102.2229
7.0273 | -102.2454
7.031 | -102.2744
7.0346 | -102.3142
7.0383 | -102.36
7.0419 | -102.4125
7.0456 | -102.4693
7.0492 | -102.5289
7.0529 | -102.5858
7.0566 | -102.6401
7.0602 | -102.6915
7.0639 | -102.7367
7.0675 | -102.7707
7.0712 | -102.787
7.0749 | -102.7897
7.0785 | -102.7815
7.0822 | -102.7616
7.0858 | -102.7253
7.0895 | -102.675
7.0931 | -102.6241
7.0968 | -102.5672
7.1005 | -102.5131
7.1041 | -102.4637
7.1078 | -102.4198
7.1114 | -102.3885
7.1151 | -102.368
7.1188 | -102.3492
7.1224 | -102.3413
7.1261 | -102.3427
7.1297 | -102.3474
7.1334 | -102.3504
7.137 | -102.3587
7.1407 | -102.3646
7.1444 | -102.3718
7.148 | -102.3786
7.1517 | -102.3842
7.1553 | -102.3864
7.159 | -102.3936
7.1627 | -102.3994
7.1663 | -102.4065
7.17 | -102.4101
7.1736 | -102.4061
7.1773 | -102.4032
7.1809 | -102.4025
7.1846 | -102.3974
7.1883 | -102.3901
7.1919 | -102.3886
7.1956 | -102.3925
7.1992 | -102.397
7.2029 | -102.3964
7.2065 | -102.3945
7.2102 | -102.395
7.2139 | -102.3976
7.2175 | -102.3954
7.2212 | -102.3887
7.2248 | -102.3832
7.2285 | -102.3738
7.2322 | -102.3678
7.2358 | -102.3636
7.2395 | -102.3648
7.2431 | -102.3665
7.2468 | -102.3725
7.2504 | -102.3809
7.2541 | -102.3864
7.2578 | -102.3912
7.2614 | -102.4016
7.2651 | -102.4127
7.2687 | -102.4253
7.2724 | -102.4386
7.2761 | -102.4595
7.2797 | -102.4868
7.2834 | -102.5133
7.287 | -102.5395
7.2907 | -102.5727
7.2943 | -102.6129
7.298 | -102.6475
7.3017 | -102.6892
7.3053 | -102.7332
7.309 | -102.7768
7.3126 | -102.8141
7.3163 | -102.8414
7.32 | -102.861
7.3236 | -102.8677
7.3273 | -102.8655
7.3309 | -102.8572
7.3346 | -102.8327
7.3382 | -102.7911
7.3419 | -102.7434
7.3456 | -102.6926
7.3492 | -102.6353
7.3529 | -102.5753
7.3565 | -102.5206
7.3602 | -102.4757
7.3638 | -102.436
7.3675 | -102.4011
7.3712 | -102.3676
7.3748 | -102.3419
7.3785 | -102.3225
7.3821 | -102.3102
7.3858 | -102.3169
7.3895 | -102.3137
7.3931 | -102.3163
7.3968 | -102.3233
7.4004 | -102.3342
7.4041 | -102.3456
7.4077 | -102.3619
7.4114 | -102.3816
7.4151 | -102.4069
7.4187 | -102.4312
7.4224 | -102.4584
7.426 | -102.4907
7.4297 | -102.5288
7.4334 | -102.5696
7.437 | -102.6047
7.4407 | -102.638
7.4443 | -102.6752
7.448 | -102.7091
7.4516 | -102.7378
7.4553 | -102.7597
7.459 | -102.7832
7.4626 | -102.8075
7.4663 | -102.8275
7.4699 | -102.8405
7.4736 | -102.8495
7.4773 | -102.8572
7.4809 | -102.8649
7.4846 | -102.8643
7.4882 | -102.8531
7.4919 | -102.8333
7.4955 | -102.8064
7.4992 | -102.778
7.5029 | -102.743
7.5065 | -102.7022
7.5102 | -102.6651
7.5138 | -102.6211
7.5175 | -102.5826
7.5211 | -102.5489
7.5248 | -102.5185
7.5285 | -102.4925
7.5321 | -102.4774
7.5358 | -102.466
7.5394 | -102.4636
7.5431 | -102.4697
7.5468 | -102.4889
7.5504 | -102.5204
7.5541 | -102.5541
7.5577 | -102.5898
7.5614 | -102.6321
7.565 | -102.6778
7.5687 | -102.7292
7.5724 | -102.7783
7.576 | -102.8265
7.5797 | -102.8741
7.5833 | -102.9182
7.587 | -102.9587
7.5907 | -102.9888
7.5943 | -103.0166
7.598 | -103.0462
7.6016 | -103.0747
7.6053 | -103.0947
7.6089 | -103.1097
7.6126 | -103.1183
7.6163 | -103.1258
7.6199 | -103.128
7.6236 | -103.122
7.6272 | -103.1153
7.6309 | -103.1012
7.6346 | -103.0836
7.6382 | -103.0624
7.6419 | -103.0366
7.6455 | -103.0102
7.6492 | -102.9819
7.6528 | -102.9478
7.6565 | -102.9127
7.6602 | -102.882
7.6638 | -102.8552
7.6675 | -102.8346
7.6711 | -102.8216
7.6748 | -102.8128
7.6784 | -102.807
7.6821 | -102.7987
7.6858 | -102.7928
7.6894 | -102.7913
7.6931 | -102.7947
7.6967 | -102.805
7.7004 | -102.8162
7.7041 | -102.8267
7.7077 | -102.8317
7.7114 | -102.8326
7.715 | -102.8383
7.7187 | -102.845
7.7223 | -102.8507
7.726 | -102.8587
7.7297 | -102.8686
7.7333 | -102.8824
7.737 | -102.8969
7.7406 | -102.9151
7.7443 | -102.934
7.748 | -102.9494
7.7516 | -102.9621
7.7553 | -102.9805
7.7589 | -102.9978
7.7626 | -103.0176
7.7662 | -103.0405
7.7699 | -103.0572
7.7736 | -103.0652
7.7772 | -103.0741
7.7809 | -103.0785
7.7845 | -103.0811
7.7882 | -103.0808
7.7919 | -103.0748
7.7955 | -103.0733
7.7992 | -103.0783
7.8028 | -103.0834
7.8065 | -103.0843
7.8101 | -103.0849
7.8138 | -103.087
7.8175 | -103.0938
7.8211 | -103.1045
7.8248 | -103.111
7.8284 | -103.1121
7.8321 | -103.1149
7.8357 | -103.1138
7.8394 | -103.1094
7.8431 | -103.1052
7.8467 | -103.1029
7.8504 | -103.105
7.854 | -103.1115
7.8577 | -103.1237
7.8614 | -103.1441
7.865 | -103.1655
7.8687 | -103.1886
7.8723 | -103.2186
7.876 | -103.2512
7.8796 | -103.2834
7.8833 | -103.3162
7.887 | -103.3461
7.8906 | -103.3652
7.8943 | -103.3848
7.8979 | -103.3979
7.9016 | -103.3974
7.9053 | -103.3933
7.9089 | -103.3785
7.9126 | -103.3521
7.9162 | -103.3168
7.9199 | -103.2888
7.9235 | -103.2656
7.9272 | -103.2395
7.9309 | -103.213
7.9345 | -103.1837
7.9382 | -103.1486
7.9418 | -103.1205
7.9455 | -103.0992
7.9492 | -103.0802
7.9528 | -103.0658
7.9565 | -103.0531
7.9601 | -103.0368
7.9638 | -103.0184
7.9674 | -103.0011
7.9711 | -102.9845
7.9748 | -102.9679
7.9784 | -102.9536
7.9821 | -102.9393
7.9857 | -102.9209
7.9894 | -102.9003
7.9931 | -102.88
7.9967 | -102.8624
8.0004 | -102.8507
8.004 | -102.8404
8.0077 | -102.8294
8.0113 | -102.8142
8.015 | -102.805
8.0187 | -102.8029
8.0223 | -102.8036
8.026 | -102.8032
8.0296 | -102.8068
8.0333 | -102.8176
8.0369 | -102.8274
8.0406 | -102.8396
8.0443 | -102.8544
8.0479 | -102.8727
8.0516 | -102.8929
8.0552 | -102.9144
8.0589 | -102.9402
8.0626 | -102.9641
8.0662 | -102.9843
8.0699 | -103.0056
8.0735 | -103.0267
8.0772 | -103.0462
8.0808 | -103.0668
8.0845 | -103.0885
8.0882 | -103.1057
8.0918 | -103.1211
8.0955 | -103.1327
8.0991 | -103.1414
8.1028 | -103.1535
8.1065 | -103.1685
8.1101 | -103.1839
8.1138 | -103.2013
8.1174 | -103.2223
8.1211 | -103.2438
8.1247 | -103.2701
8.1284 | -103.3018
8.1321 | -103.3362
8.1357 | -103.3605
8.1394 | -103.3811
8.143 | -103.3931
8.1467 | -103.4049
8.1503 | -103.4155
8.154 | -103.4121
8.1577 | -103.409
8.1613 | -103.4074
8.165 | -103.4021
8.1686 | -103.3955
8.1723 | -103.3873
8.176 | -103.3754
8.1796 | -103.3609
8.1833 | -103.3444
8.1869 | -103.3303
8.1906 | -103.3119
8.1942 | -103.2963
8.1979 | -103.2806
8.2016 | -103.2678
8.2052 | -103.2515
8.2089 | -103.2306
8.2125 | -103.2157
8.2162 | -103.2053
8.2199 | -103.1953
8.2235 | -103.1849
8.2272 | -103.1755
8.2308 | -103.1719
8.2345 | -103.1731
8.2381 | -103.1787
8.2418 | -103.1863
8.2455 | -103.1983
8.2491 | -103.2154
8.2528 | -103.2317
8.2564 | -103.2522
8.2601 | -103.2793
8.2638 | -103.3106
8.2674 | -103.3373
8.2711 | -103.3626
8.2747 | -103.3908
8.2784 | -103.4191
8.282 | -103.4417
8.2857 | -103.4568
8.2894 | -103.4652
8.293 | -103.468
8.2967 | -103.4701
8.3003 | -103.4677
8.304 | -103.4575
8.3077 | -103.4381
8.3113 | -103.4231
8.315 | -103.4056
8.3186 | -103.3924
8.3223 | -103.3781
8.3259 | -103.3701
8.3296 | -103.3608
8.3333 | -103.3516
8.3369 | -103.344
8.3406 | -103.3341
8.3442 | -103.3236
8.3479 | -103.3125
8.3515 | -103.2968
8.3552 | -103.2833
8.3589 | -103.2685
8.3625 | -103.2525
8.3662 | -103.2382
8.3698 | -103.2302
8.3735 | -103.2332
8.3772 | -103.2454
8.3808 | -103.2583
8.3845 | -103.274
8.3881 | -103.302
8.3918 | -103.3346
8.3954 | -103.3801
8.3991 | -103.4235
8.4028 | -103.4714
8.4064 | -103.5277
8.4101 | -103.5856
8.4137 | -103.6423
8.4174 | -103.6884
8.4211 | -103.7282
8.4247 | -103.767
8.4284 | -103.8015
8.432 | -103.8224
8.4357 | -103.8334
8.4393 | -103.8297
8.443 | -103.8033
8.4467 | -103.7843
8.4503 | -103.7637
8.454 | -103.7481
8.4576 | -103.7207
8.4613 | -103.6909
8.4649 | -103.6659
8.4686 | -103.6401
8.4723 | -103.6126
8.4759 | -103.5839
8.4796 | -103.5513
8.4832 | -103.5258
8.4869 | -103.5046
8.4906 | -103.4845
8.4942 | -103.4631
8.4979 | -103.4419
8.5015 | -103.4217
8.5052 | -103.4099
8.5088 | -103.4059
8.5125 | -103.4044
8.5162 | -103.4101
8.5198 | -103.4169
8.5235 | -103.4333
8.5271 | -103.451
8.5308 | -103.4684
8.5345 | -103.4901
8.5381 | -103.5118
8.5418 | -103.5339
8.5454 | -103.5463
8.5491 | -103.5539
8.5527 | -103.5565
8.5564 | -103.5554
8.5601 | -103.551
8.5637 | -103.5441
8.5674 | -103.532
8.571 | -103.5195
8.5747 | -103.5071
8.5784 | -103.4902
8.582 | -103.4683
8.5857 | -103.4486
8.5893 | -103.4311
8.593 | -103.4156
8.5966 | -103.4065
8.6003 | -103.4029
8.604 | -103.4022
8.6076 | -103.4067
8.6113 | -103.4153
8.6149 | -103.4322
8.6186 | -103.4537
8.6223 | -103.4763
8.6259 | -103.5023
8.6296 | -103.5371
8.6332 | -103.5723
8.6369 | -103.6082
8.6405 | -103.642
8.6442 | -103.6694
8.6479 | -103.6911
8.6515 | -103.7115
8.6552 | -103.7329
8.6588 | -103.748
8.6625 | -103.7537
8.6661 | -103.7506
8.6698 | -103.7422
8.6735 | -103.7352
8.6771 | -103.7223
8.6808 | -103.706
8.6844 | -103.6896
8.6881 | -103.6685
8.6918 | -103.6414
8.6954 | -103.6162
8.6991 | -103.5963
8.7027 | -103.5797
8.7064 | -103.5621
8.71 | -103.5484
8.7137 | -103.5378
8.7174 | -103.5339
8.721 | -103.5328
8.7247 | -103.5324
8.7283 | -103.5404
8.732 | -103.5562
8.7357 | -103.5755
8.7393 | -103.5929
8.743 | -103.6139
8.7466 | -103.6403
8.7503 | -103.6712
8.7539 | -103.7023
8.7576 | -103.7269
8.7613 | -103.7518
8.7649 | -103.777
8.7686 | -103.802
8.7722 | -103.8263
8.7759 | -103.851
8.7796 | -103.8643
8.7832 | -103.8672
8.7869 | -103.8688
8.7905 | -103.8682
8.7942 | -103.8709
8.7978 | -103.8703
8.8015 | -103.8632
8.8052 | -103.8513
8.8088 | -103.8429
8.8125 | -103.8301
8.8161 | -103.8229
8.8198 | -103.8219
8.8234 | -103.8149
8.8271 | -103.8043
8.8308 | -103.795
8.8344 | -103.791
8.8381 | -103.7917
8.8417 | -103.7977
8.8454 | -103.7997
8.8491 | -103.805
8.8527 | -103.8111
8.8564 | -103.8136
8.86 | -103.8203
8.8637 | -103.8251
8.8673 | -103.8325
8.871 | -103.8384
8.8747 | -103.8439
8.8783 | -103.8543
8.882 | -103.8651
8.8856 | -103.8734
8.8893 | -103.8813
8.893 | -103.8868
8.8966 | -103.8893
8.9003 | -103.8964
8.9039 | -103.903
8.9076 | -103.9153
8.9112 | -103.9307
8.9149 | -103.9436
8.9186 | -103.9575
8.9222 | -103.9689
8.9259 | -103.9856
8.9295 | -104.0084
8.9332 | -104.0334
8.9369 | -104.0583
8.9405 | -104.09
8.9442 | -104.1231
8.9478 | -104.1568
8.9515 | -104.1933
8.9551 | -104.2292
8.9588 | -104.2649
8.9625 | -104.2966
8.9661 | -104.3263
8.9698 | -104.3478
8.9734 | -104.3639
8.9771 | -104.3758
8.9807 | -104.3847
8.9844 | -104.3937
8.9881 | -104.3931
8.9917 | -104.3826
8.9954 | -104.3664
8.999 | -104.3421
9.0027 | -104.3089
9.0064 | -104.2679
9.01 | -104.2209
9.0137 | -104.1748
9.0173 | -104.1323
9.021 | -104.0874
9.0246 | -104.048
9.0283 | -104.0148
9.032 | -103.9801
9.0356 | -103.9528
9.0393 | -103.9329
9.0429 | -103.9197
9.0466 | -103.9118
9.0503 | -103.9109
9.0539 | -103.9202
9.0576 | -103.9306
9.0612 | -103.9484
9.0649 | -103.9677
9.0685 | -103.9846
9.0722 | -104.0002
9.0759 | -104.0142
9.0795 | -104.0269
9.0832 | -104.0391
9.0868 | -104.0454
9.0905 | -104.0433
9.0942 | -104.0385
9.0978 | -104.0292
9.1015 | -104.02
9.1051 | -104.0063
9.1088 | -103.995
9.1124 | -103.9868
9.1161 | -103.9763
9.1198 | -103.9647
9.1234 | -103.9599
9.1271 | -103.9555
9.1307 | -103.9569
9.1344 | -103.9596
9.138 | -103.9626
9.1417 | -103.9689
9.1454 | -103.9793
9.149 | -103.9868
9.1527 | -103.9917
9.1563 | -104.0018
9.16 | -104.011
9.1637 | -104.0176
9.1673 | -104.0246
9.171 | -104.0258
9.1746 | -104.0245
9.1783 | -104.0215
9.1819 | -104.0139
9.1856 | -104.0027
9.1893 | -103.9876
9.1929 | -103.9705
9.1966 | -103.9481
9.2002 | -103.9314
9.2039 | -103.9151
9.2076 | -103.8945
9.2112 | -103.8785
9.2149 | -103.8635
9.2185 | -103.8524
9.2222 | -103.8476
9.2258 | -103.847
9.2295 | -103.8445
9.2332 | -103.8435
9.2368 | -103.8398
9.2405 | -103.8458
9.2441 | -103.8514
9.2478 | -103.8554
9.2515 | -103.8627
9.2551 | -103.871
9.2588 | -103.8797
9.2624 | -103.8923
9.2661 | -103.9032
9.2697 | -103.9176
9.2734 | -103.9358
9.2771 | -103.9491
9.2807 | -103.9643
9.2844 | -103.9852
9.288 | -104.0092
9.2917 | -104.025
9.2953 | -104.0404
9.299 | -104.0599
9.3027 | -104.0787
9.3063 | -104.0989
9.31 | -104.1195
9.3136 | -104.137
9.3173 | -104.1606
9.321 | -104.181
9.3246 | -104.1965
9.3283 | -104.2113
9.3319 | -104.2205
9.3356 | -104.2245
9.3392 | -104.2274
9.3429 | -104.2321
9.3466 | -104.2281
9.3502 | -104.2205
9.3539 | -104.209
9.3575 | -104.199
9.3612 | -104.1881
9.3649 | -104.1797
9.3685 | -104.1673
9.3722 | -104.1583
9.3758 | -104.1459
9.3795 | -104.137
9.3831 | -104.1348
9.3868 | -104.1301
9.3905 | -104.1271
9.3941 | -104.1269
9.3978 | -104.1262
9.4014 | -104.1253
9.4051 | -104.1204
9.4088 | -104.1168
9.4124 | -104.1165
9.4161 | -104.1119
9.4197 | -104.1029
9.4234 | -104.0975
9.427 | -104.0963
9.4307 | -104.0927
9.4344 | -104.0863
9.438 | -104.079
9.4417 | -104.0746
9.4453 | -104.0691
9.449 | -104.0687
9.4527 | -104.0661
9.4563 | -104.0668
9.46 | -104.0717
9.4636 | -104.0785
9.4673 | -104.086
9.4709 | -104.0899
9.4746 | -104.0937
9.4783 | -104.0958
9.4819 | -104.098
9.4856 | -104.0955
9.4892 | -104.0935
9.4929 | -104.0908
9.4965 | -104.0778
9.5002 | -104.0746
9.5039 | -104.0658
9.5075 | -104.0563
9.5112 | -104.0449
9.5148 | -104.0328
9.5185 | -104.0228
9.5222 | -104.0121
9.5258 | -104.0025
9.5295 | -103.9967
9.5331 | -103.9955
9.5368 | -103.9996
9.5404 | -104.0089
9.5441 | -104.0235
9.5478 | -104.0378
9.5514 | -104.0532
9.5551 | -104.0713
9.5587 | -104.0953
9.5624 | -104.1193
9.5661 | -104.147
9.5697 | -104.1746
9.5734 | -104.2016
9.577 | -104.2284
9.5807 | -104.2512
9.5843 | -104.2781
9.588 | -104.3057
9.5917 | -104.3348
9.5953 | -104.3634
9.599 | -104.3886
9.6026 | -104.4097
9.6063 | -104.4345
9.6099 | -104.4596
9.6136 | -104.4778
9.6173 | -104.4953
9.6209 | -104.5107
9.6246 | -104.5263
9.6282 | -104.5326
9.6319 | -104.5353
9.6356 | -104.5378
9.6392 | -104.5355
9.6429 | -104.5309
9.6465 | -104.5279
9.6502 | -104.5223
9.6538 | -104.5111
9.6575 | -104.5002
9.6612 | -104.4873
9.6648 | -104.4753
9.6685 | -104.4667
9.6721 | -104.4564
9.6758 | -104.4456
9.6795 | -104.4369
9.6831 | -104.4317
9.6868 | -104.4249
9.6904 | -104.4267
9.6941 | -104.4325
9.6977 | -104.44
9.7014 | -104.4474
9.7051 | -104.4525
9.7087 | -104.4538
9.7124 | -104.4559
9.716 | -104.4612
9.7197 | -104.4646
9.7234 | -104.4607
9.727 | -104.4488
9.7307 | -104.4349
9.7343 | -104.4206
9.738 | -104.4013
9.7416 | -104.3791
9.7453 | -104.3575
9.749 | -104.3313
9.7526 | -104.3055
9.7563 | -104.285
9.7599 | -104.2594
9.7636 | -104.2364
9.7673 | -104.2199
9.7709 | -104.2098
9.7746 | -104.2013
9.7782 | -104.198
9.7819 | -104.1992
9.7855 | -104.2012
9.7892 | -104.2015
9.7929 | -104.2084
9.7965 | -104.2226
9.8002 | -104.2379
9.8038 | -104.2544
9.8075 | -104.2694
9.8111 | -104.287
9.8148 | -104.3045
9.8185 | -104.3194
9.8221 | -104.3328
9.8258 | -104.3417
9.8294 | -104.3478
9.8331 | -104.3446
9.8368 | -104.3433
9.8404 | -104.3394
9.8441 | -104.3312
9.8477 | -104.3257
9.8514 | -104.3219
9.855 | -104.3172
9.8587 | -104.3081
9.8624 | -104.3051
9.866 | -104.3034
9.8697 | -104.3012
9.8733 | -104.2997
9.877 | -104.3028
9.8807 | -104.314
9.8843 | -104.3321
9.888 | -104.35
9.8916 | -104.3666
9.8953 | -104.3839
9.8989 | -104.4038
9.9026 | -104.4275
9.9063 | -104.4514
9.9099 | -104.476
9.9136 | -104.5013
9.9172 | -104.5261
9.9209 | -104.5525
9.9246 | -104.5812
9.9282 | -104.6076
9.9319 | -104.6369
9.9355 | -104.6603
9.9392 | -104.6801
9.9428 | -104.6952
9.9465 | -104.7083
9.9502 | -104.7227
9.9538 | -104.733
9.9575 | -104.7417
9.9611 | -104.7494
9.9648 | -104.7523
9.9684 | -104.7517
9.9721 | -104.7472
9.9758 | -104.7417
9.9794 | -104.7286
9.9831 | -104.7113
9.9867 | -104.6962
9.9904 | -104.6772
9.9941 | -104.6548
9.9977 | -104.6313
10.0014 | -104.6077
10.005 | -104.5841
10.0087 | -104.5659
10.0123 | -104.5517
10.016 | -104.5371
10.0197 | -104.5252
10.0233 | -104.5134
10.027 | -104.525
10.0306 | -104.5244
10.0343 | -104.5248
10.038 | -104.5313
10.0416 | -104.5424
10.0453 | -104.5527
10.0489 | -104.5623
10.0526 | -104.5729
10.0562 | -104.5806
10.0599 | -104.5836
10.0636 | -104.5889
10.0672 | -104.596
10.0709 | -104.5999
10.0745 | -104.6037
10.0782 | -104.6075
10.0819 | -104.6092
10.0855 | -104.6099
10.0892 | -104.6114
10.0928 | -104.6098
10.0965 | -104.6031
10.1001 | -104.5998
10.1038 | -104.5982
10.1075 | -104.5945
10.1111 | -104.594
10.1148 | -104.5975
10.1184 | -104.6001
10.1221 | -104.6025
10.1257 | -104.6046
10.1294 | -104.6107
10.1331 | -104.6168
10.1367 | -104.6206
10.1404 | -104.6291
10.144 | -104.635
10.1477 | -104.6413
10.1514 | -104.6467
10.155 | -104.6489
10.1587 | -104.6471
10.1623 | -104.6435
10.166 | -104.6352
10.1696 | -104.6261
10.1733 | -104.6127
10.177 | -104.5935
10.1806 | -104.572
10.1843 | -104.5502
10.1879 | -104.5222
10.1916 | -104.496
10.1953 | -104.4722
10.1989 | -104.4501
10.2026 | -104.4287
10.2062 | -104.4107
10.2099 | -104.4002
10.2135 | -104.3881
10.2172 | -104.3707
10.2209 | -104.3587
10.2245 | -104.3543
10.2282 | -104.3518
10.2318 | -104.3542
10.2355 | -104.3602
10.2392 | -104.3704
10.2428 | -104.3858
10.2465 | -104.4032
10.2501 | -104.4192
10.2538 | -104.4335
10.2574 | -104.4527
10.2611 | -104.474
10.2648 | -104.4953
10.2684 | -104.5187
10.2721 | -104.5369
10.2757 | -104.5585
10.2794 | -104.58
10.283 | -104.5989
10.2867 | -104.6199
10.2904 | -104.6359
10.294 | -104.6519
10.2977 | -104.6578
10.3013 | -104.6622
10.305 | -104.6713
10.3087 | -104.6805
10.3123 | -104.691
10.316 | -104.6959
10.3196 | -104.6988
10.3233 | -104.7025
10.3269 | -104.7042
10.3306 | -104.7075
10.3343 | -104.709
10.3379 | -104.7069
10.3416 | -104.6996
10.3452 | -104.692
10.3489 | -104.6846
10.3526 | -104.6756
10.3562 | -104.6653
10.3599 | -104.653
10.3635 | -104.6425
10.3672 | -104.6355
10.3708 | -104.6247
10.3745 | -104.6097
10.3782 | -104.5955
10.3818 | -104.5869
10.3855 | -104.5795
10.3891 | -104.5726
10.3928 | -104.5694
10.3965 | -104.5692
10.4001 | -104.572
10.4038 | -104.5785
10.4074 | -104.5844
10.4111 | -104.5935
10.4147 | -104.6036
10.4184 | -104.6129
10.4221 | -104.6237
10.4257 | -104.6316
10.4294 | -104.639
10.433 | -104.6469
10.4367 | -104.6514
10.4403 | -104.6569
10.444 | -104.6608
10.4477 | -104.6688
10.4513 | -104.6709
10.455 | -104.6678
10.4586 | -104.6603
10.4623 | -104.651
10.466 | -104.6473
10.4696 | -104.6428
10.4733 | -104.6342
10.4769 | -104.6308
10.4806 | -104.6245
10.4842 | -104.6195
10.4879 | -104.6128
10.4916 | -104.6055
10.4952 | -104.6005
10.4989 | -104.5973
10.5025 | -104.5942
10.5062 | -104.5966
10.5099 | -104.603
10.5135 | -104.6094
10.5172 | -104.6184
10.5208 | -104.6313
10.5245 | -104.6424
10.5281 | -104.6578
10.5318 | -104.6717
10.5355 | -104.6895
10.5391 | -104.7102
10.5428 | -104.7291
10.5464 | -104.7519
10.5501 | -104.7715
10.5538 | -104.789
10.5574 | -104.8058
10.5611 | -104.8189
10.5647 | -104.8275
10.5684 | -104.8394
10.572 | -104.8536
10.5757 | -104.8676
10.5794 | -104.8799
10.583 | -104.889
10.5867 | -104.894
10.5903 | -104.9029
10.594 | -104.9105
10.5976 | -104.9175
10.6013 | -104.922
10.605 | -104.9219
10.6086 | -104.9219
10.6123 | -104.9279
10.6159 | -104.9285
10.6196 | -104.9247
10.6233 | -104.9227
10.6269 | -104.9185
10.6306 | -104.9081
10.6342 | -104.897
10.6379 | -104.8877
10.6415 | -104.8768
10.6452 | -104.8571
10.6489 | -104.8344
10.6525 | -104.8132
10.6562 | -104.7935
10.6598 | -104.7696
10.6635 | -104.7451
10.6672 | -104.7257
10.6708 | -104.7086
10.6745 | -104.6931
10.6781 | -104.6827
10.6818 | -104.6689
10.6854 | -104.6537
10.6891 | -104.641
10.6928 | -104.6336
10.6964 | -104.6252
10.7001 | -104.6201
10.7037 | -104.6205
10.7074 | -104.6218
10.7111 | -104.6288
10.7147 | -104.6378
10.7184 | -104.6505
10.722 | -104.6608
10.7257 | -104.6649
10.7293 | -104.6702
10.733 | -104.6761
10.7367 | -104.6802
10.7403 | -104.6847
10.744 | -104.6933
10.7476 | -104.6952
10.7513 | -104.6931
10.7549 | -104.6937
10.7586 | -104.6954
10.7623 | -104.7014
10.7659 | -104.7061
10.7696 | -104.7027
10.7732 | -104.7029
10.7769 | -104.7079
10.7806 | -104.7134
10.7842 | -104.721
10.7879 | -104.7289
10.7915 | -104.7396
10.7952 | -104.7497
10.7988 | -104.7611
10.8025 | -104.7781
10.8062 | -104.7954
10.8098 | -104.8125
10.8135 | -104.8305
10.8171 | -104.8478
10.8208 | -104.8634
10.8245 | -104.8693
10.8281 | -104.8727
10.8318 | -104.8799
10.8354 | -104.8846
10.8391 | -104.888
10.8427 | -104.8883
10.8464 | -104.884
10.8501 | -104.8718
10.8537 | -104.8588
10.8574 | -104.8385
10.861 | -104.8136
10.8647 | -104.7891
10.8684 | -104.7582
10.872 | -104.7256
10.8757 | -104.6922
10.8793 | -104.6602
10.883 | -104.6324
10.8866 | -104.6035
10.8903 | -104.5713
10.894 | -104.5387
10.8976 | -104.5065
10.9013 | -104.4731
10.9049 | -104.4416
10.9086 | -104.4169
10.9123 | -104.3943
10.9159 | -104.3707
10.9196 | -104.3513
10.9232 | -104.3334
10.9269 | -104.3201
10.9305 | -104.3103
10.9342 | -104.3055
10.9379 | -104.3026
10.9415 | -104.3015
10.9452 | -104.3089
10.9488 | -104.3196
10.9525 | -104.3322
10.9561 | -104.3504
10.9598 | -104.3689
10.9635 | -104.3906
10.9671 | -104.4153
10.9708 | -104.443
10.9744 | -104.4738
10.9781 | -104.5072
10.9818 | -104.5476
10.9854 | -104.588
10.9891 | -104.6232
10.9927 | -104.66
10.9964 | -104.698
11 | -104.7379
11.0037 | -104.7757
11.0074 | -104.8132
11.011 | -104.8491
11.0147 | -104.8818
11.0183 | -104.9098
11.022 | -104.9282
11.0257 | -104.9492
11.0293 | -104.9656
11.033 | -104.973
11.0366 | -104.9757
11.0403 | -104.9725
11.0439 | -104.964
11.0476 | -104.9532
11.0513 | -104.9393
11.0549 | -104.9237
11.0586 | -104.9092
11.0622 | -104.8967
11.0659 | -104.8814
11.0696 | -104.8659
11.0732 | -104.8574
11.0769 | -104.8504
11.0805 | -104.8727
11.0842 | -104.8804
11.0878 | -104.8953
11.0915 | -104.9099
11.0952 | -104.921
11.0988 | -104.9339
11.1025 | -104.9523
11.1061 | -104.968
11.1098 | -104.9869
11.1134 | -105.0065
11.1171 | -105.024
11.1208 | -105.0398
11.1244 | -105.0518
11.1281 | -105.0619
11.1317 | -105.0689
11.1354 | -105.0751
11.1391 | -105.0793
11.1427 | -105.0796
11.1464 | -105.0767
11.15 | -105.0751
11.1537 | -105.0725
11.1573 | -105.0654
11.161 | -105.0588
11.1647 | -105.0566
11.1683 | -105.0556
11.172 | -105.0545
11.1756 | -105.0497
11.1793 | -105.0478
11.183 | -105.0515
11.1866 | -105.0509
11.1903 | -105.0511
11.1939 | -105.0569
11.1976 | -105.0641
11.2012 | -105.0721
11.2049 | -105.0819
11.2086 | -105.094
11.2122 | -105.1067
11.2159 | -105.1199
11.2195 | -105.1276
11.2232 | -105.1316
11.2269 | -105.1363
11.2305 | -105.1395
11.2342 | -105.1401
11.2378 | -105.1364
11.2415 | -105.123
11.2451 | -105.1065
11.2488 | -105.0896
11.2525 | -105.0707
11.2561 | -105.0524
11.2598 | -105.0332
11.2634 | -105.0153
11.2671 | -104.9943
11.2707 | -104.9702
11.2744 | -104.9409
11.2781 | -104.9173
11.2817 | -104.8975
11.2854 | -104.8804
11.289 | -104.8657
11.2927 | -104.8531
11.2964 | -104.8433
11.3 | -104.8375
11.3037 | -104.8327
11.3073 | -104.8303
11.311 | -104.8317
11.3146 | -104.8319
11.3183 | -104.8358
11.322 | -104.8441
11.3256 | -104.8508
11.3293 | -104.8556
11.3329 | -104.8622
11.3366 | -104.8729
11.3403 | -104.8874
11.3439 | -104.9017
11.3476 | -104.9171
11.3512 | -104.9349
11.3549 | -104.9525
11.3585 | -104.9711
11.3622 | -104.9882
11.3659 | -105.0067
11.3695 | -105.0246
11.3732 | -105.0432
11.3768 | -105.0631
11.3805 | -105.0774
11.3842 | -105.0923
11.3878 | -105.1084
11.3915 | -105.1216
11.3951 | -105.1331
11.3988 | -105.142
11.4024 | -105.1494
11.4061 | -105.1568
11.4098 | -105.1631
11.4134 | -105.1695
11.4171 | -105.1747
11.4207 | -105.1766
11.4244 | -105.1798
11.428 | -105.1789
11.4317 | -105.1773
11.4354 | -105.1799
11.439 | -105.1825
11.4427 | -105.179
11.4463 | -105.1734
11.45 | -105.1669
11.4537 | -105.1613
11.4573 | -105.1573
11.461 | -105.1498
11.4646 | -105.1384
11.4683 | -105.1291
11.4719 | -105.1194
11.4756 | -105.1077
11.4793 | -105.093
11.4829 | -105.0768
11.4866 | -105.0632
11.4902 | -105.0504
11.4939 | -105.0355
11.4976 | -105.019
11.5012 | -105.0041
11.5049 | -104.9921
11.5085 | -104.9792
11.5122 | -104.9617
11.5158 | -104.944
11.5195 | -104.9282
11.5232 | -104.9132
11.5268 | -104.9032
11.5305 | -104.8865
11.5341 | -104.8704
11.5378 | -104.8567
11.5415 | -104.8415
11.5451 | -104.8319
11.5488 | -104.823
11.5524 | -104.8135
11.5561 | -104.8077
11.5597 | -104.8024
11.5634 | -104.8027
11.5671 | -104.8054
11.5707 | -104.8105
11.5744 | -104.8152
11.578 | -104.8202
11.5817 | -104.8268
11.5853 | -104.8374
11.589 | -104.8509
11.5927 | -104.8693
11.5963 | -104.8873
11.6 | -104.9017
11.6036 | -104.9184
11.6073 | -104.9366
11.611 | -104.9509
11.6146 | -104.9643
11.6183 | -104.9762
11.6219 | -104.987
11.6256 | -104.9953
11.6292 | -105.0044
11.6329 | -105.0129
11.6366 | -105.0223
11.6402 | -105.0321
11.6439 | -105.0415
11.6475 | -105.0531
11.6512 | -105.0618
11.6549 | -105.0658
11.6585 | -105.0719
11.6622 | -105.0808
11.6658 | -105.09
11.6695 | -105.0945
11.6731 | -105.0964
11.6768 | -105.0998
11.6805 | -105.1062
11.6841 | -105.1095
11.6878 | -105.1145
11.6914 | -105.1169
11.6951 | -105.1173
11.6988 | -105.1212
11.7024 | -105.1235
11.7061 | -105.1222
11.7097 | -105.1198
11.7134 | -105.1158
11.717 | -105.1146
11.7207 | -105.1109
11.7244 | -105.105
11.728 | -105.1004
11.7317 | -105.093
11.7353 | -105.086
11.739 | -105.0814
11.7426 | -105.081
11.7463 | -105.0776
11.75 | -105.0704
11.7536 | -105.0632
11.7573 | -105.0551
11.7609 | -105.0467
11.7646 | -105.0412
11.7683 | -105.0362
11.7719 | -105.0335
11.7756 | -105.0342
11.7792 | -105.0377
11.7829 | -105.0409
11.7865 | -105.0455
11.7902 | -105.0526
11.7939 | -105.0573
11.7975 | -105.0661
11.8012 | -105.0766
11.8048 | -105.0906
11.8085 | -105.101
11.8122 | -105.1108
11.8158 | -105.1211
11.8195 | -105.1321
11.8231 | -105.1432
11.8268 | -105.1537
11.8304 | -105.1582
11.8341 | -105.1575
11.8378 | -105.1564
11.8414 | -105.1528
11.8451 | -105.1469
11.8487 | -105.141
11.8524 | -105.1335
11.8561 | -105.1234
11.8597 | -105.1075
11.8634 | -105.0906
11.867 | -105.0731
11.8707 | -105.0575
11.8743 | -105.0419
11.878 | -105.0266
11.8817 | -105.0123
11.8853 | -104.9944
11.889 | -104.9743
11.8926 | -104.9588
11.8963 | -104.9453
11.9 | -104.9343
11.9036 | -104.9279
11.9073 | -104.9242
11.9109 | -104.9202
11.9146 | -104.9188
11.9182 | -104.9219
11.9219 | -104.9297
11.9256 | -104.9394
11.9292 | -104.9506
11.9329 | -104.9665
11.9365 | -104.9831
11.9402 | -105.0036
11.9438 | -105.0252
11.9475 | -105.0434
11.9512 | -105.0613
11.9548 | -105.0835
11.9585 | -105.1023
11.9621 | -105.1234
11.9658 | -105.1447
11.9695 | -105.1626
11.9731 | -105.1777
11.9768 | -105.1883
11.9804 | -105.1993
11.9841 | -105.2062
11.9877 | -105.2133
11.9914 | -105.2198
11.9951 | -105.2249
11.9987 | -105.2287
12.0024 | -105.2321
12.006 | -105.237
12.0097 | -105.2367
12.0134 | -105.237
12.017 | -105.2364
12.0207 | -105.2286
12.0243 | -105.2219
12.028 | -105.2193
12.0316 | -105.2194
12.0353 | -105.2214
12.039 | -105.2248
12.0426 | -105.2278
12.0463 | -105.2302
12.0499 | -105.233
12.0536 | -105.2397
12.0572 | -105.2526
12.0609 | -105.2646
12.0646 | -105.2754
12.0682 | -105.2843
12.0719 | -105.2916
12.0755 | -105.2998
12.0792 | -105.3086
12.0829 | -105.3171
12.0865 | -105.3252
12.0902 | -105.3362
12.0938 | -105.3484
12.0975 | -105.3592
12.1011 | -105.3674
12.1048 | -105.3713
12.1085 | -105.3785
12.1121 | -105.3862
12.1158 | -105.3932
12.1194 | -105.403
12.1231 | -105.4109
12.1268 | -105.4178
12.1304 | -105.427
12.1341 | -105.4313
12.1377 | -105.4334
12.1414 | -105.4349
12.145 | -105.4352
12.1487 | -105.4383
12.1524 | -105.4409
12.156 | -105.44
12.1597 | -105.4402
12.1633 | -105.4371
12.167 | -105.4355
12.1707 | -105.4354
12.1743 | -105.4336
12.178 | -105.4332
12.1816 | -105.429
12.1853 | -105.4239
12.1889 | -105.4174
12.1926 | -105.4119
12.1963 | -105.4109
12.1999 | -105.4075
12.2036 | -105.4007
12.2072 | -105.392
12.2109 | -105.3816
12.2145 | -105.3687
12.2182 | -105.3587
12.2219 | -105.3529
12.2255 | -105.3451
12.2292 | -105.336
12.2328 | -105.3237
12.2365 | -105.3118
12.2402 | -105.3047
12.2438 | -105.3001
12.2475 | -105.296
12.2511 | -105.2905
12.2548 | -105.2854
12.2584 | -105.2829
12.2621 | -105.281
12.2658 | -105.2786
12.2694 | -105.2791
12.2731 | -105.2801
12.2767 | -105.2848
12.2804 | -105.2905
12.2841 | -105.2979
12.2877 | -105.3078
12.2914 | -105.3186
12.295 | -105.3315
12.2987 | -105.3502
12.3023 | -105.3684
12.306 | -105.3828
12.3097 | -105.3999
12.3133 | -105.4192
12.317 | -105.4367
12.3206 | -105.4526
12.3243 | -105.4705
12.328 | -105.4852
12.3316 | -105.5018
12.3353 | -105.5195
12.3389 | -105.5339
12.3426 | -105.5438
12.3462 | -105.5477
12.3499 | -105.5488
12.3536 | -105.5496
12.3572 | -105.5496
12.3609 | -105.5533
12.3645 | -105.5563
12.3682 | -105.5557
12.3719 | -105.5515
12.3755 | -105.5451
12.3792 | -105.538
12.3828 | -105.5311
12.3865 | -105.524
12.3901 | -105.5152
12.3938 | -105.5074
12.3975 | -105.4979
12.4011 | -105.488
12.4048 | -105.4788
12.4084 | -105.4675
12.4121 | -105.4617
12.4157 | -105.4599
12.4194 | -105.4534
12.4231 | -105.4472
12.4267 | -105.4439
12.4304 | -105.4439
12.434 | -105.4475
12.4377 | -105.4504
12.4414 | -105.4534
12.445 | -105.4598
12.4487 | -105.4651
12.4523 | -105.4744
12.456 | -105.4823
12.4596 | -105.488
12.4633 | -105.4925
12.467 | -105.4968
12.4706 | -105.5026
12.4743 | -105.5064
12.4779 | -105.5089
12.4816 | -105.5124
12.4853 | -105.5181
12.4889 | -105.5233
12.4926 | -105.5283
12.4962 | -105.5367
12.4999 | -105.5463
12.5035 | -105.5531
12.5072 | -105.5584
12.5109 | -105.5648
12.5145 | -105.5703
12.5182 | -105.5761
12.5218 | -105.5848
12.5255 | -105.5936
12.5292 | -105.6021
12.5328 | -105.6125
12.5365 | -105.618
12.5401 | -105.621
12.5438 | -105.6231
12.5474 | -105.6255
12.5511 | -105.6296
12.5548 | -105.6327
12.5584 | -105.6354
12.5621 | -105.636
12.5657 | -105.6341
12.5694 | -105.6347
12.573 | -105.6329
12.5767 | -105.6308
12.5804 | -105.6287
12.584 | -105.6272
12.5877 | -105.6262
12.5913 | -105.6196
12.595 | -105.6114
12.5987 | -105.6048
12.6023 | -105.6002
12.606 | -105.5962
12.6096 | -105.5914
12.6133 | -105.5844
12.6169 | -105.5805
12.6206 | -105.5777
12.6243 | -105.5731
12.6279 | -105.5688
12.6316 | -105.5632
12.6352 | -105.5575
12.6389 | -105.5549
12.6426 | -105.5537
12.6462 | -105.5514
12.6499 | -105.5506
12.6535 | -105.5542
12.6572 | -105.5554
12.6608 | -105.5555
12.6645 | -105.5646
12.6682 | -105.5693
12.6718 | -105.5717
12.6755 | -105.5774
12.6791 | -105.5851
12.6828 | -105.5934
12.6865 | -105.6025
12.6901 | -105.6144
12.6938 | -105.6206
12.6974 | -105.6286
12.7011 | -105.638
12.7047 | -105.6451
12.7084 | -105.6474
12.7121 | -105.6491
12.7157 | -105.6508
12.7194 | -105.6518
12.723 | -105.65
12.7267 | -105.6488
12.7303 | -105.6456
12.734 | -105.6373
12.7377 | -105.6273
12.7413 | -105.6174
12.745 | -105.6089
12.7486 | -105.5986
12.7523 | -105.5808
12.756 | -105.5637
12.7596 | -105.5445
12.7633 | -105.5269
12.7669 | -105.5101
12.7706 | -105.493
12.7742 | -105.4797
12.7779 | -105.467
12.7816 | -105.4532
12.7852 | -105.4401
12.7889 | -105.429
12.7925 | -105.4204
12.7962 | -105.4146
12.7999 | -105.4097
12.8035 | -105.4076
12.8072 | -105.4102
12.8108 | -105.4161
12.8145 | -105.4227
12.8181 | -105.4319
12.8218 | -105.4435
12.8255 | -105.4553
12.8291 | -105.4687
12.8328 | -105.4828
12.8364 | -105.4975
12.8401 | -105.5164
12.8438 | -105.5342
12.8474 | -105.55
12.8511 | -105.5659
12.8547 | -105.5801
12.8584 | -105.5948
12.862 | -105.6071
12.8657 | -105.6174
12.8694 | -105.6258
12.873 | -105.6335
12.8767 | -105.64
12.8803 | -105.6458
12.884 | -105.6482
12.8876 | -105.6465
12.8913 | -105.6411
12.895 | -105.6358
12.8986 | -105.6266
12.9023 | -105.6129
12.9059 | -105.5963
12.9096 | -105.5777
12.9133 | -105.557
12.9169 | -105.5352
12.9206 | -105.5096
12.9242 | -105.4828
12.9279 | -105.4555
12.9315 | -105.4292
12.9352 | -105.4064
12.9389 | -105.3827
12.9425 | -105.3565
12.9462 | -105.3306
12.9498 | -105.3043
12.9535 | -105.2764
12.9572 | -105.2546
12.9608 | -105.2319
12.9645 | -105.21
12.9681 | -105.1892
12.9718 | -105.1669
12.9754 | -105.1451
12.9791 | -105.1256
12.9828 | -105.1093
12.9864 | -105.096
12.9901 | -105.0859
12.9937 | -105.0744
12.9974 | -105.0653
13.0011 | -105.0591
13.0047 | -105.0548
13.0084 | -105.0478
13.012 | -105.045
13.0157 | -105.0498
13.0193 | -105.0568
13.023 | -105.0636
13.0267 | -105.0719
13.0303 | -105.0861
13.034 | -105.1007
13.0376 | -105.1112
13.0413 | -105.1227
13.0449 | -105.1349
13.0486 | -105.1463
13.0523 | -105.1563
13.0559 | -105.1652
13.0596 | -105.1731
13.0632 | -105.1786
13.0669 | -105.186
13.0706 | -105.1923
13.0742 | -105.1974
13.0779 | -105.1996
13.0815 | -105.2
13.0852 | -105.1978
13.0888 | -105.1957
13.0925 | -105.1935
13.0962 | -105.1919
13.0998 | -105.1895
13.1035 | -105.1861
13.1071 | -105.1864
13.1108 | -105.1889
13.1145 | -105.1911
13.1181 | -105.1941
13.1218 | -105.1983
13.1254 | -105.2096
13.1291 | -105.2233
13.1327 | -105.2371
13.1364 | -105.2521
13.1401 | -105.2732
13.1437 | -105.2935
13.1474 | -105.3158
13.151 | -105.3437
13.1547 | -105.3753
13.1584 | -105.408
13.162 | -105.441
13.1657 | -105.4725
13.1693 | -105.504
13.173 | -105.5354
13.1766 | -105.5657
13.1803 | -105.5971
13.184 | -105.6278
13.1876 | -105.6586
13.1913 | -105.6754
13.1949 | -105.6986
13.1986 | -105.7235
13.2022 | -105.7429
13.2059 | -105.7566
13.2096 | -105.7685
13.2132 | -105.7827
13.2169 | -105.7927
13.2205 | -105.7983
13.2242 | -105.802
13.2279 | -105.8056
13.2315 | -105.8058
13.2352 | -105.8013
13.2388 | -105.7941
13.2425 | -105.7845
13.2461 | -105.7693
13.2498 | -105.7493
13.2535 | -105.7286
13.2571 | -105.7018
13.2608 | -105.671
13.2644 | -105.6382
13.2681 | -105.6058
13.2718 | -105.5699
13.2754 | -105.5349
13.2791 | -105.4944
13.2827 | -105.4533
13.2864 | -105.4137
13.29 | -105.3767
13.2937 | -105.3401
13.2974 | -105.303
13.301 | -105.2631
13.3047 | -105.2239
13.3083 | -105.1884
13.312 | -105.1563
13.3157 | -105.1275
13.3193 | -105.1009
13.323 | -105.0792
13.3266 | -105.0631
13.3303 | -105.0495
13.3339 | -105.0386
13.3376 | -105.0331
13.3413 | -105.0283
13.3449 | -105.0279
13.3486 | -105.029
13.3522 | -105.0334
13.3559 | -105.0405
13.3596 | -105.0502
13.3632 | -105.0613
13.3669 | -105.0712
13.3705 | -105.0823
13.3742 | -105.0946
13.3778 | -105.1084
13.3815 | -105.1216
13.3852 | -105.1348
13.3888 | -105.1513
13.3925 | -105.1637
13.3961 | -105.1709
13.3998 | -105.1774
13.4034 | -105.1846
13.4071 | -105.1875
13.4108 | -105.1895
13.4144 | -105.1905
13.4181 | -105.1933
13.4217 | -105.196
13.4254 | -105.1982
13.4291 | -105.1974
13.4327 | -105.1946
13.4364 | -105.1939
13.44 | -105.1944
13.4437 | -105.1929
13.4473 | -105.1883
13.451 | -105.1865
13.4547 | -105.1867
13.4583 | -105.1887
13.462 | -105.1939
13.4656 | -105.2014
13.4693 | -105.2054
13.473 | -105.212
13.4766 | -105.2202
13.4803 | -105.2276
13.4839 | -105.237
13.4876 | -105.2525
13.4912 | -105.2673
13.4949 | -105.2859
13.4986 | -105.3086
13.5022 | -105.3279
13.5059 | -105.3472
13.5095 | -105.3659
13.5132 | -105.3883
13.5168 | -105.4105
13.5205 | -105.4278
13.5242 | -105.4437
13.5278 | -105.4589
13.5315 | -105.4769
13.5351 | -105.4949
13.5388 | -105.512
13.5425 | -105.531
13.5461 | -105.5474
13.5498 | -105.5612
13.5534 | -105.5702
13.5571 | -105.5776
13.5607 | -105.5863
13.5644 | -105.5956
13.5681 | -105.6015
13.5717 | -105.6029
13.5754 | -105.6021
13.579 | -105.6045
13.5827 | -105.6019
13.5864 | -105.6
13.59 | -105.6005
13.5937 | -105.5971
13.5973 | -105.5906
13.601 | -105.5839
13.6046 | -105.5775
13.6083 | -105.5688
13.612 | -105.5553
13.6156 | -105.5408
13.6193 | -105.5256
13.6229 | -105.512
13.6266 | -105.4949
13.6303 | -105.4735
13.6339 | -105.4519
13.6376 | -105.4361
13.6412 | -105.4193
13.6449 | -105.3969
13.6485 | -105.3745
13.6522 | -105.3525
13.6559 | -105.3322
13.6595 | -105.3149
13.6632 | -105.2997
13.6668 | -105.2863
13.6705 | -105.2745
13.6742 | -105.2639
13.6778 | -105.254
13.6815 | -105.2435
13.6851 | -105.2389
13.6888 | -105.2345
13.6924 | -105.2333
13.6961 | -105.2385
13.6998 | -105.2474
13.7034 | -105.2589
13.7071 | -105.2702
13.7107 | -105.2858
13.7144 | -105.307
13.718 | -105.3403
13.7217 | -105.3649
13.7254 | -105.3943
13.729 | -105.4201
13.7327 | -105.4471
13.7363 | -105.4763
13.74 | -105.5062
13.7437 | -105.5381
13.7473 | -105.5712
13.751 | -105.6041
13.7546 | -105.6353
13.7583 | -105.6653
13.7619 | -105.6897
13.7656 | -105.7138
13.7693 | -105.7351
13.7729 | -105.7553
13.7766 | -105.7689
13.7802 | -105.7781
13.7839 | -105.7877
13.7876 | -105.7972
13.7912 | -105.8052
13.7949 | -105.8072
13.7985 | -105.803
13.8022 | -105.7963
13.8058 | -105.7909
13.8095 | -105.7881
13.8132 | -105.7833
13.8168 | -105.7762
13.8205 | -105.7664
13.8241 | -105.7589
13.8278 | -105.7512
13.8315 | -105.7422
13.8351 | -105.7294
13.8388 | -105.7181
13.8424 | -105.7103
13.8461 | -105.7003
13.8497 | -105.6855
13.8534 | -105.6675
13.8571 | -105.6509
13.8607 | -105.6402
13.8644 | -105.6307
13.868 | -105.6211
13.8717 | -105.6111
13.8753 | -105.5987
13.879 | -105.5846
13.8827 | -105.5701
13.8863 | -105.5599
13.89 | -105.5457
13.8936 | -105.531
13.8973 | -105.5146
13.901 | -105.4963
13.9046 | -105.4805
13.9083 | -105.4648
13.9119 | -105.4473
13.9156 | -105.4287
13.9192 | -105.4145
13.9229 | -105.4009
13.9266 | -105.3843
13.9302 | -105.3695
13.9339 | -105.3577
13.9375 | -105.3442
13.9412 | -105.3324
13.9449 | -105.3226
13.9485 | -105.3149
13.9522 | -105.3108
13.9558 | -105.3061
13.9595 | -105.2886
13.9631 | -105.269
13.9668 | -105.2493
13.9705 | -105.2387
13.9741 | -105.2289
13.9778 | -105.2183
13.9814 | -105.208
13.9851 | -105.1947
13.9888 | -105.1789
13.9924 | -105.1642
13.9961 | -105.1516
13.9997 | -105.1393
14.0034 | -105.1245
14.007 | -105.1052
14.0107 | -105.081
14.0144 | -105.0573
14.018 | -105.0326
14.0217 | -105.0071
14.0253 | -104.9831
14.029 | -104.9582
14.0326 | -104.9313
14.0363 | -104.9048
14.04 | -104.8819
14.0436 | -104.862
14.0473 | -104.8451
14.0509 | -104.8337
14.0546 | -104.825
14.0583 | -104.8191
14.0619 | -104.8152
14.0656 | -104.8159
14.0692 | -104.8215
14.0729 | -104.832
14.0765 | -104.8494
14.0802 | -104.8736
14.0839 | -104.9012
14.0875 | -104.9306
14.0912 | -104.963
14.0948 | -104.9964
14.0985 | -105.0317
14.1022 | -105.0719
14.1058 | -105.1109
14.1095 | -105.149
14.1131 | -105.1889
14.1168 | -105.2288
14.1204 | -105.2645
14.1241 | -105.2997
14.1278 | -105.3347
14.1314 | -105.3692
14.1351 | -105.3995
14.1387 | -105.429
14.1424 | -105.4536
14.1461 | -105.4763
14.1497 | -105.4974
14.1534 | -105.5131
14.157 | -105.5282
14.1607 | -105.5504
14.1643 | -105.5708
14.168 | -105.5821
14.1717 | -105.5823
14.1753 | -105.5765
14.179 | -105.5644
14.1826 | -105.555
14.1863 | -105.5443
14.1899 | -105.5309
14.1936 | -105.5129
14.1973 | -105.4916
14.2009 | -105.469
14.2046 | -105.4461
14.2082 | -105.4226
14.2119 | -105.3968
14.2156 | -105.371
14.2192 | -105.3449
14.2229 | -105.3186
14.2265 | -105.2924
14.2302 | -105.2686
14.2338 | -105.2475
14.2375 | -105.2258
14.2412 | -105.2019
14.2448 | -105.1788
14.2485 | -105.1645
14.2521 | -105.1445
14.2558 | -105.1267
14.2595 | -105.1117
14.2631 | -105.0971
14.2668 | -105.0853
14.2704 | -105.076
14.2741 | -105.0695
14.2777 | -105.066
14.2814 | -105.0593
14.2851 | -105.0542
14.2887 | -105.0511
14.2924 | -105.0481
14.296 | -105.0471
14.2997 | -105.0501
14.3034 | -105.0529
14.307 | -105.0565
14.3107 | -105.0603
14.3143 | -105.063
14.318 | -105.0644
14.3216 | -105.0668
14.3253 | -105.0699
14.329 | -105.0709
14.3326 | -105.0762
14.3363 | -105.0792
14.3399 | -105.0829
14.3436 | -105.0868
14.3473 | -105.0906
14.3509 | -105.0928
14.3546 | -105.0943
14.3582 | -105.0952
14.3619 | -105.0965
14.3655 | -105.0988
14.3692 | -105.1006
14.3729 | -105.1042
14.3765 | -105.1058
14.3802 | -105.1096
14.3838 | -105.1153
14.3875 | -105.1162
14.3911 | -105.1198
14.3948 | -105.123
14.3985 | -105.1264
14.4021 | -105.1333
14.4058 | -105.1378
14.4094 | -105.1414
14.4131 | -105.1471
14.4168 | -105.156
14.4204 | -105.1635
14.4241 | -105.1691
14.4277 | -105.1743
14.4314 | -105.1776
14.435 | -105.1789
14.4387 | -105.1812
14.4424 | -105.1849
14.446 | -105.1892
14.4497 | -105.1913
14.4533 | -105.1932
14.457 | -105.1969
14.4607 | -105.199
14.4643 | -105.2007
14.468 | -105.2002
14.4716 | -105.2001
14.4753 | -105.1966
14.4789 | -105.1917
14.4826 | -105.1868
14.4863 | -105.1806
14.4899 | -105.1773
14.4936 | -105.1743
14.4972 | -105.1694
14.5009 | -105.1661
14.5046 | -105.1628
14.5082 | -105.157
14.5119 | -105.1525
14.5155 | -105.1478
14.5192 | -105.1436
14.5228 | -105.1403
14.5265 | -105.1368
14.5302 | -105.1322
14.5338 | -105.1271
14.5375 | -105.1178
14.5411 | -105.1057
14.5448 | -105.0972
14.5484 | -105.0895
14.5521 | -105.0838
14.5558 | -105.0798
14.5594 | -105.0762
14.5631 | -105.0722
14.5667 | -105.0683
14.5704 | -105.0636
14.5741 | -105.0564
14.5777 | -105.0476
14.5814 | -105.0398
14.585 | -105.0312
14.5887 | -105.0209
14.5923 | -105.0127
14.596 | -105.0058
14.5997 | -104.9966
14.6033 | -104.9895
14.607 | -104.9813
14.6106 | -104.9719
14.6143 | -104.9659
14.618 | -104.9582
14.6216 | -104.9501
14.6253 | -104.9412
14.6289 | -104.9359
14.6326 | -104.9343
14.6362 | -104.9321
14.6399 | -104.928
14.6436 | -104.9265
14.6472 | -104.9276
14.6509 | -104.9257
14.6545 | -104.926
14.6582 | -104.9289
14.6618 | -104.9332
14.6655 | -104.9375
14.6692 | -104.9412
14.6728 | -104.9449
14.6765 | -104.9504
14.6801 | -104.9565
14.6838 | -104.9646
14.6875 | -104.9739
14.6911 | -104.983
14.6948 | -104.9927
14.6984 | -105.0048
14.7021 | -105.0168
14.7057 | -105.0291
14.7094 | -105.0405
14.7131 | -105.0496
14.7167 | -105.0613
14.7204 | -105.0727
14.724 | -105.0841
14.7277 | -105.0997
14.7314 | -105.1147
14.735 | -105.1321
14.7387 | -105.15
14.7423 | -105.1665
14.746 | -105.1855
14.7496 | -105.2033
14.7533 | -105.2175
14.757 | -105.2299
14.7606 | -105.2403
14.7643 | -105.2502
14.7679 | -105.2565
14.7716 | -105.2651
14.7753 | -105.2741
14.7789 | -105.2818
14.7826 | -105.2847
14.7862 | -105.2909
14.7899 | -105.2951
14.7935 | -105.2965
14.7972 | -105.2992
14.8009 | -105.302
14.8045 | -105.3011
14.8082 | -105.2978
14.8118 | -105.2961
14.8155 | -105.2926
14.8192 | -105.2889
14.8228 | -105.2828
14.8265 | -105.276
14.8301 | -105.2679
14.8338 | -105.2589
14.8374 | -105.2516
14.8411 | -105.2422
14.8448 | -105.2325
14.8484 | -105.2238
14.8521 | -105.215
14.8557 | -105.2071
14.8594 | -105.2008
14.863 | -105.1955
14.8667 | -105.1866
14.8704 | -105.1782
14.874 | -105.1731
14.8777 | -105.1703
14.8813 | -105.1702
14.885 | -105.1708
14.8887 | -105.1727
14.8923 | -105.1749
14.896 | -105.1783
14.8996 | -105.1841
14.9033 | -105.1894
14.9069 | -105.1918
14.9106 | -105.1973
14.9143 | -105.2033
14.9179 | -105.209
14.9216 | -105.2158
14.9252 | -105.2268
14.9289 | -105.2364
14.9326 | -105.2461
14.9362 | -105.2562
14.9399 | -105.2616
14.9435 | -105.2664
14.9472 | -105.2699
14.9508 | -105.2728
14.9545 | -105.2772
14.9582 | -105.2803
14.9618 | -105.2816
14.9655 | -105.2815
14.9691 | -105.2778
14.9728 | -105.2745
14.9765 | -105.2715
14.9801 | -105.267
14.9838 | -105.257
14.9874 | -105.2417
14.9911 | -105.2282
14.9947 | -105.2179
14.9984 | -105.208
15.0021 | -105.193
15.0057 | -105.1763
15.0094 | -105.1594
15.013 | -105.1434
15.0167 | -105.1276
15.0203 | -105.1132
15.024 | -105.0998
15.0277 | -105.0865
15.0313 | -105.0739
15.035 | -105.0636
15.0386 | -105.057
15.0423 | -105.0499
15.046 | -105.043
15.0496 | -105.0351
15.0533 | -105.0266
15.0569 | -105.0216
15.0606 | -105.0186
15.0642 | -105.0187
15.0679 | -105.0235
15.0716 | -105.0284
15.0752 | -105.033
15.0789 | -105.0402
15.0825 | -105.0494
15.0862 | -105.0587
15.0899 | -105.0658
15.0935 | -105.0762
15.0972 | -105.0851
15.1008 | -105.0936
15.1045 | -105.1017
15.1081 | -105.1087
15.1118 | -105.1144
15.1155 | -105.1207
15.1191 | -105.1295
15.1228 | -105.1395
15.1264 | -105.1442
15.1301 | -105.147
15.1338 | -105.1512
15.1374 | -105.1538
15.1411 | -105.1542
15.1447 | -105.1522
15.1484 | -105.1499
15.152 | -105.1445
15.1557 | -105.1403
15.1594 | -105.1394
15.163 | -105.1377
15.1667 | -105.135
15.1703 | -105.1303
15.174 | -105.1259
15.1776 | -105.1199
15.1813 | -105.115
15.185 | -105.1117
15.1886 | -105.1079
15.1923 | -105.1019
15.1959 | -105.0962
15.1996 | -105.0932
15.2033 | -105.0906
15.2069 | -105.0869
15.2106 | -105.0834
15.2142 | -105.0817
15.2179 | -105.0784
15.2215 | -105.0767
15.2252 | -105.0765
15.2289 | -105.0755
15.2325 | -105.0762
15.2362 | -105.0772
15.2398 | -105.0804
15.2435 | -105.0828
15.2472 | -105.0871
15.2508 | -105.089
15.2545 | -105.0913
15.2581 | -105.0981
15.2618 | -105.1065
15.2654 | -105.1157
15.2691 | -105.1242
15.2728 | -105.13
15.2764 | -105.1382
15.2801 | -105.1438
15.2837 | -105.1463
15.2874 | -105.1487
15.2911 | -105.1509
15.2947 | -105.1539
15.2984 | -105.1557
15.302 | -105.1596
15.3057 | -105.162
15.3093 | -105.1655
15.313 | -105.1673
15.3167 | -105.1709
15.3203 | -105.1732
15.324 | -105.1745
15.3276 | -105.1793
15.3313 | -105.1834
15.3349 | -105.1831
15.3386 | -105.1804
15.3423 | -105.1779
15.3459 | -105.1774
15.3496 | -105.1779
15.3532 | -105.1766
15.3569 | -105.1747
15.3606 | -105.1706
15.3642 | -105.167
15.3679 | -105.1634
15.3715 | -105.1593
15.3752 | -105.1525
15.3788 | -105.1438
15.3825 | -105.1361
15.3862 | -105.1297
15.3898 | -105.1254
15.3935 | -105.1198
15.3971 | -105.1141
15.4008 | -105.1102
15.4045 | -105.1041
15.4081 | -105.0964
15.4118 | -105.0918
15.4154 | -105.0873
15.4191 | -105.0837
15.4227 | -105.083
15.4264 | -105.0794
15.4301 | -105.0748
15.4337 | -105.0721
15.4374 | -105.0714
15.441 | -105.0704
15.4447 | -105.0698
15.4484 | -105.0666
15.452 | -105.0636
15.4557 | -105.0638
15.4593 | -105.0628
15.463 | -105.0589
15.4666 | -105.055
15.4703 | -105.0541
15.474 | -105.0507
15.4776 | -105.0433
15.4813 | -105.0376
15.4849 | -105.0292
15.4886 | -105.0212
15.4922 | -105.0131
15.4959 | -105.0047
15.4996 | -104.9968
15.5032 | -104.9883
15.5069 | -104.9798
15.5105 | -104.9707
15.5142 | -104.9624
15.5179 | -104.9535
15.5215 | -104.9395
15.5252 | -104.9245
15.5288 | -104.9111
15.5325 | -104.8974
15.5361 | -104.8815
15.5398 | -104.8654
15.5435 | -104.8513
15.5471 | -104.8382
15.5508 | -104.8253
15.5544 | -104.8126
15.5581 | -104.8007
15.5618 | -104.7898
15.5654 | -104.7804
15.5691 | -104.7682
15.5727 | -104.7562
15.5764 | -104.7481
15.58 | -104.7397
15.5837 | -104.734
15.5874 | -104.7283
15.591 | -104.7222
15.5947 | -104.718
15.5983 | -104.7152
15.602 | -104.7129
15.6057 | -104.712
15.6093 | -104.7089
15.613 | -104.7055
15.6166 | -104.7051
15.6203 | -104.7041
15.6239 | -104.7024
15.6276 | -104.7048
15.6313 | -104.7078
15.6349 | -104.7107
15.6386 | -104.716
15.6422 | -104.7207
15.6459 | -104.7239
15.6495 | -104.7308
15.6532 | -104.7387
15.6569 | -104.7458
15.6605 | -104.7542
15.6642 | -104.7649
15.6678 | -104.7745
15.6715 | -104.7862
15.6752 | -104.7995
15.6788 | -104.8132
15.6825 | -104.8248
15.6861 | -104.8402
15.6898 | -104.8564
15.6934 | -104.8682
15.6971 | -104.8796
15.7008 | -104.8935
15.7044 | -104.907
15.7081 | -104.9211
15.7117 | -104.932
15.7154 | -104.9418
15.7191 | -104.9477
15.7227 | -104.953
15.7264 | -104.9572
15.73 | -104.9568
15.7337 | -104.9543
15.7373 | -104.9476
15.741 | -104.9412
15.7447 | -104.9338
15.7483 | -104.9252
15.752 | -104.9162
15.7556 | -104.9037
15.7593 | -104.8905
15.763 | -104.8762
15.7666 | -104.8598
15.7703 | -104.8446
15.7739 | -104.8268
15.7776 | -104.808
15.7812 | -104.7903
15.7849 | -104.7733
15.7886 | -104.7567
15.7922 | -104.7424
15.7959 | -104.7295
15.7995 | -104.7168
15.8032 | -104.7058
15.8069 | -104.6946
15.8105 | -104.6856
15.8142 | -104.6792
15.8178 | -104.6747
15.8215 | -104.6728
15.8251 | -104.6719
15.8288 | -104.6695
15.8325 | -104.6796
15.8361 | -104.6834
15.8398 | -104.6892
15.8434 | -104.6955
15.8471 | -104.7046
15.8507 | -104.7149
15.8544 | -104.7257
15.8581 | -104.7397
15.8617 | -104.7547
15.8654 | -104.7691
15.869 | -104.7832
15.8727 | -104.7988
15.8764 | -104.8147
15.88 | -104.8311
15.8837 | -104.8472
15.8873 | -104.8661
15.891 | -104.8851
15.8946 | -104.9047
15.8983 | -104.925
15.902 | -104.9428
15.9056 | -104.9598
15.9093 | -104.981
15.9129 | -105.0008
15.9166 | -105.0176
15.9203 | -105.0379
15.9239 | -105.0566
15.9276 | -105.0713
15.9312 | -105.0853
15.9349 | -105.0998
15.9385 | -105.1172
15.9422 | -105.1337
15.9459 | -105.1478
15.9495 | -105.1624
15.9532 | -105.1783
15.9568 | -105.1951
15.9605 | -105.2118
15.9642 | -105.2304
15.9678 | -105.247
15.9715 | -105.2633
15.9751 | -105.2797
15.9788 | -105.2951
15.9824 | -105.3126
15.9861 | -105.3295
15.9898 | -105.3446
15.9934 | -105.3592
15.9971 | -105.372
16.0007 | -105.3829
16.0044 | -105.3957
16.008 | -105.407
16.0117 | -105.4159
16.0154 | -105.4242
16.019 | -105.4307
16.0227 | -105.4354
16.0263 | -105.436
16.03 | -105.4369
16.0337 | -105.4361
16.0373 | -105.4321
16.041 | -105.4251
16.0446 | -105.4143
16.0483 | -105.403
16.0519 | -105.3893
16.0556 | -105.375
16.0593 | -105.361
16.0629 | -105.3455
16.0666 | -105.327
16.0702 | -105.305
16.0739 | -105.2809
16.0776 | -105.2552
16.0812 | -105.2281
16.0849 | -105.202
16.0885 | -105.1744
16.0922 | -105.1476
16.0958 | -105.1243
16.0995 | -105.0992
16.1032 | -105.0739
16.1068 | -105.0496
16.1105 | -105.0268
16.1141 | -105.0053
16.1178 | -104.9846
16.1214 | -104.9634
16.1251 | -104.9438
16.1288 | -104.9258
16.1324 | -104.9101
16.1361 | -104.8967
16.1397 | -104.8859
16.1434 | -104.8754
16.1471 | -104.8679
16.1507 | -104.8603
16.1544 | -104.8547
16.158 | -104.851
16.1617 | -104.8492
16.1653 | -104.849
16.169 | -104.8484
16.1727 | -104.8482
16.1763 | -104.8518
16.18 | -104.8568
16.1836 | -104.8613
16.1873 | -104.8664
16.191 | -104.8731
16.1946 | -104.8802
16.1983 | -104.8881
16.2019 | -104.8958
16.2056 | -104.9038
16.2092 | -104.9114
16.2129 | -104.9183
16.2166 | -104.9279
16.2202 | -104.9409
16.2239 | -104.9537
16.2275 | -104.9644
16.2312 | -104.9764
16.2349 | -104.9892
16.2385 | -105.0023
16.2422 | -105.014
16.2458 | -105.0257
16.2495 | -105.0381
16.2531 | -105.051
16.2568 | -105.0653
16.2605 | -105.0786
16.2641 | -105.0942
16.2678 | -105.1113
16.2714 | -105.127
16.2751 | -105.1424
16.2788 | -105.157
16.2824 | -105.1724
16.2861 | -105.1895
16.2897 | -105.2048
16.2934 | -105.2175
16.297 | -105.2297
16.3007 | -105.2453
16.3044 | -105.2629
16.308 | -105.2791
16.3117 | -105.294
16.3153 | -105.306
16.319 | -105.319
16.3226 | -105.3308
16.3263 | -105.342
16.33 | -105.3566
16.3336 | -105.3684
16.3373 | -105.3747
16.3409 | -105.3794
16.3446 | -105.3859
16.3483 | -105.3903
16.3519 | -105.3946
16.3556 | -105.3965
16.3592 | -105.3952
16.3629 | -105.3947
16.3665 | -105.3929
16.3702 | -105.3905
16.3739 | -105.387
16.3775 | -105.3867
16.3812 | -105.3878
16.3848 | -105.3875
16.3885 | -105.3869
16.3922 | -105.3883
16.3958 | -105.3872
16.3995 | -105.3855
16.4031 | -105.3802
16.4068 | -105.3765
16.4104 | -105.3714
16.4141 | -105.3663
16.4178 | -105.3619
16.4214 | -105.3578
16.4251 | -105.3522
16.4287 | -105.3426
16.4324 | -105.3344
16.4361 | -105.3265
16.4397 | -105.3179
16.4434 | -105.3093
16.447 | -105.2996
16.4507 | -105.2859
16.4543 | -105.2725
16.458 | -105.259
16.4617 | -105.2458
16.4653 | -105.2326
16.469 | -105.2175
16.4726 | -105.2025
16.4763 | -105.1896
16.4799 | -105.1728
16.4836 | -105.1539
16.4873 | -105.1371
16.4909 | -105.1213
16.4946 | -105.1045
16.4982 | -105.0881
16.5019 | -105.0728
16.5056 | -105.0606
16.5092 | -105.0487
16.5129 | -105.0352
16.5165 | -105.0242
16.5202 | -105.0126
16.5238 | -104.9997
16.5275 | -104.9917
16.5312 | -104.9857
16.5348 | -104.9786
16.5385 | -104.9716
16.5421 | -104.969
16.5458 | -104.9668
16.5495 | -104.9659
16.5531 | -104.9626
16.5568 | -104.9621
16.5604 | -104.9631
16.5641 | -104.9635
16.5677 | -104.9673
16.5714 | -104.9735
16.5751 | -104.9808
16.5787 | -104.9874
16.5824 | -104.9937
16.586 | -104.9988
16.5897 | -105.0066
16.5934 | -105.0117
16.597 | -105.0173
16.6007 | -105.0256
16.6043 | -105.0351
16.608 | -105.0425
16.6116 | -105.0477
16.6153 | -105.0497
16.619 | -105.0506
16.6226 | -105.0536
16.6263 | -105.0555
16.6299 | -105.0548
16.6336 | -105.0546
16.6372 | -105.0559
16.6409 | -105.0589
16.6446 | -105.059
16.6482 | -105.0592
16.6519 | -105.0585
16.6555 | -105.0572
16.6592 | -105.0544
16.6629 | -105.0503
16.6665 | -105.0485
16.6702 | -105.0451
16.6738 | -105.0414
16.6775 | -105.0372
16.6811 | -105.0307
16.6848 | -105.0253
16.6885 | -105.0209
16.6921 | -105.0143
16.6958 | -105.0098
16.6994 | -105.0055
16.7031 | -105.0021
16.7068 | -104.9973
16.7104 | -104.9922
16.7141 | -104.9891
16.7177 | -104.9882
16.7214 | -104.9883
16.725 | -104.9842
16.7287 | -104.9805
16.7324 | -104.977
16.736 | -104.9743
16.7397 | -104.9724
16.7433 | -104.9707
16.747 | -104.9707
16.7507 | -104.9728
16.7543 | -104.974
16.758 | -104.9757
16.7616 | -104.9777
16.7653 | -104.978
16.7689 | -104.98
16.7726 | -104.982
16.7763 | -104.9828
16.7799 | -104.9817
16.7836 | -104.9846
16.7872 | -104.9889
16.7909 | -104.9953
16.7945 | -105.0015
16.7982 | -105.0076
16.8019 | -105.0148
16.8055 | -105.0222
16.8092 | -105.0296
16.8128 | -105.0375
16.8165 | -105.0459
16.8202 | -105.0548
16.8238 | -105.068
16.8275 | -105.0784
16.8311 | -105.0895
16.8348 | -105.1002
16.8384 | -105.1108
16.8421 | -105.1189
16.8458 | -105.1261
16.8494 | -105.1354
16.8531 | -105.1463
16.8567 | -105.154
16.8604 | -105.159
16.8641 | -105.1656
16.8677 | -105.1721
16.8714 | -105.1777
16.875 | -105.1829
16.8787 | -105.1884
16.8823 | -105.1917
16.886 | -105.1837
16.8897 | -105.1843
16.8933 | -105.1844
16.897 | -105.1845
16.9006 | -105.1839
16.9043 | -105.1853
16.908 | -105.1863
16.9116 | -105.1848
16.9153 | -105.1816
16.9189 | -105.1809
16.9226 | -105.1828
16.9262 | -105.1848
16.9299 | -105.1861
16.9336 | -105.1845
16.9372 | -105.1835
16.9409 | -105.1815
16.9445 | -105.1793
16.9482 | -105.1802
16.9519 | -105.1823
16.9555 | -105.1824
16.9592 | -105.1819
16.9628 | -105.1813
16.9665 | -105.1827
16.9701 | -105.1848
16.9738 | -105.1887
16.9775 | -105.1904
16.9811 | -105.1904
16.9848 | -105.1911
16.9884 | -105.1937
16.9921 | -105.1953
16.9957 | -105.1979
16.9994 | -105.2025
17.0031 | -105.2079
17.0067 | -105.2112
17.0104 | -105.213
17.014 | -105.2171
17.0177 | -105.2225
17.0214 | -105.2261
17.025 | -105.2262
17.0287 | -105.2279
17.0323 | -105.2313
17.036 | -105.2324
17.0396 | -105.2311
17.0433 | -105.2293
17.047 | -105.2276
17.0506 | -105.2262
17.0543 | -105.2226
17.0579 | -105.2208
17.0616 | -105.2183
17.0653 | -105.214
17.0689 | -105.2112
17.0726 | -105.208
17.0762 | -105.2032
17.0799 | -105.195
17.0835 | -105.1869
17.0872 | -105.1809
17.0909 | -105.1756
17.0945 | -105.1697
17.0982 | -105.1645
17.1018 | -105.1604
17.1055 | -105.1545
17.1091 | -105.1487
17.1128 | -105.1438
17.1165 | -105.1398
17.1201 | -105.135
17.1238 | -105.1331
17.1274 | -105.1342
17.1311 | -105.1332
17.1348 | -105.1315
17.1384 | -105.1286
17.1421 | -105.1244
17.1457 | -105.1232
17.1494 | -105.1188
17.153 | -105.1154
17.1567 | -105.1143
17.1604 | -105.1158
17.164 | -105.114
17.1677 | -105.111
17.1713 | -105.1067
17.175 | -105.105
17.1787 | -105.1053
17.1823 | -105.1033
17.186 | -105.1029
17.1896 | -105.1034
17.1933 | -105.1032
17.1969 | -105.1026
17.2006 | -105.1015
17.2043 | -105.1007
17.2079 | -105.1027
17.2116 | -105.1023
17.2152 | -105.1
17.2189 | -105.0967
17.2226 | -105.0959
17.2262 | -105.0969
17.2299 | -105.0974
17.2335 | -105.0941
17.2372 | -105.0912
17.2408 | -105.0879
17.2445 | -105.0848
17.2482 | -105.0831
17.2518 | -105.0823
17.2555 | -105.0815
17.2591 | -105.078
17.2628 | -105.0717
17.2664 | -105.0686
17.2701 | -105.066
17.2738 | -105.0607
17.2774 | -105.0564
17.2811 | -105.0528
17.2847 | -105.0502
17.2884 | -105.0466
17.2921 | -105.0399
17.2957 | -105.0352
17.2994 | -105.0309
17.303 | -105.0221
17.3067 | -105.0157
17.3103 | -105.0108
17.314 | -105.0078
17.3177 | -105.0016
17.3213 | -104.996
17.325 | -104.995
17.3286 | -104.9954
17.3323 | -104.9932
17.336 | -104.9876
17.3396 | -104.9829
17.3433 | -104.9803
17.3469 | -104.9795
17.3506 | -104.9792
17.3542 | -104.9796
17.3579 | -104.9805
17.3616 | -104.9808
17.3652 | -104.9832
17.3689 | -104.985
17.3725 | -104.9882
17.3762 | -104.9926
17.3799 | -104.997
17.3835 | -105
17.3872 | -105.0047
17.3908 | -105.0086
17.3945 | -105.0133
17.3981 | -105.0201
17.4018 | -105.0274
17.4055 | -105.0327
17.4091 | -105.0358
17.4128 | -105.0339
17.4164 | -105.0353
17.4201 | -105.0399
17.4238 | -105.0418
17.4274 | -105.0429
17.4311 | -105.0456
17.4347 | -105.0465
17.4384 | -105.0428
17.442 | -105.0393
17.4457 | -105.038
17.4494 | -105.0368
17.453 | -105.0334
17.4567 | -105.028
17.4603 | -105.0233
17.464 | -105.019
17.4676 | -105.0152
17.4713 | -105.008
17.475 | -104.9986
17.4786 | -104.9888
17.4823 | -104.9799
17.4859 | -104.973
17.4896 | -104.9635
17.4933 | -104.9516
17.4969 | -104.9387
17.5006 | -104.9266
17.5042 | -104.9153
17.5079 | -104.906
17.5115 | -104.897
17.5152 | -104.8862
17.5189 | -104.8755
17.5225 | -104.8659
17.5262 | -104.8564
17.5298 | -104.8459
17.5335 | -104.8356
17.5372 | -104.8276
17.5408 | -104.8214
17.5445 | -104.8115
17.5481 | -104.8023
17.5518 | -104.7947
17.5554 | -104.7877
17.5591 | -104.7796
17.5628 | -104.7704
17.5664 | -104.7668
17.5701 | -104.7626
17.5737 | -104.7565
17.5774 | -104.7493
17.5811 | -104.7449
17.5847 | -104.7434
17.5884 | -104.7419
17.592 | -104.7394
17.5957 | -104.7353
17.5993 | -104.7323
17.603 | -104.7298
17.6067 | -104.7275
17.6103 | -104.7269
17.614 | -104.726
17.6176 | -104.7269
17.6213 | -104.7271
17.6249 | -104.7277
17.6286 | -104.7277
17.6323 | -104.7306
17.6359 | -104.7344
17.6396 | -104.7374
17.6432 | -104.7402
17.6469 | -104.7386
17.6506 | -104.7373
17.6542 | -104.7376
17.6579 | -104.7413
17.6615 | -104.7431
17.6652 | -104.744
17.6688 | -104.7471
17.6725 | -104.7475
17.6762 | -104.749
17.6798 | -104.7492
17.6835 | -104.7491
17.6871 | -104.7489
17.6908 | -104.7488
17.6945 | -104.748
17.6981 | -104.747
17.7018 | -104.7472
17.7054 | -104.7486
17.7091 | -104.7476
17.7127 | -104.7475
17.7164 | -104.7464
17.7201 | -104.7454
17.7237 | -104.7451
17.7274 | -104.7462
17.731 | -104.7478
17.7347 | -104.7492
17.7384 | -104.7452
17.742 | -104.743
17.7457 | -104.741
17.7493 | -104.7402
17.753 | -104.739
17.7566 | -104.7377
17.7603 | -104.734
17.764 | -104.7322
17.7676 | -104.7294
17.7713 | -104.7266
17.7749 | -104.7233
17.7786 | -104.7186
17.7822 | -104.7129
17.7859 | -104.7079
17.7896 | -104.702
17.7932 | -104.696
17.7969 | -104.6923
17.8005 | -104.6875
17.8042 | -104.6841
17.8079 | -104.6804
17.8115 | -104.6765
17.8152 | -104.6719
17.8188 | -104.6658
17.8225 | -104.6612
17.8261 | -104.6546
17.8298 | -104.6473
17.8335 | -104.6429
17.8371 | -104.6396
17.8408 | -104.6341
17.8444 | -104.6285
17.8481 | -104.6259
17.8518 | -104.6248
17.8554 | -104.6229
17.8591 | -104.6205
17.8627 | -104.6163
17.8664 | -104.611
17.87 | -104.6064
17.8737 | -104.6019
17.8774 | -104.597
17.881 | -104.5934
17.8847 | -104.5887
17.8883 | -104.583
17.892 | -104.5779
17.8957 | -104.574
17.8993 | -104.5709
17.903 | -104.5667
17.9066 | -104.56
17.9103 | -104.5544
17.9139 | -104.551
17.9176 | -104.5465
17.9213 | -104.5394
17.9249 | -104.5333
17.9286 | -104.5263
17.9322 | -104.5194
17.9359 | -104.5138
17.9395 | -104.509
17.9432 | -104.5088
17.9469 | -104.5075
17.9505 | -104.5056
17.9542 | -104.5029
17.9578 | -104.5009
17.9615 | -104.4994
17.9652 | -104.4989
17.9688 | -104.4999
17.9725 | -104.5007
17.9761 | -104.5017
17.9798 | -104.5038
17.9834 | -104.5065
17.9871 | -104.5082
17.9908 | -104.5097
17.9944 | -104.5142
17.9981 | -104.5205
18.0017 | -104.5252
18.0054 | -104.5301
18.0091 | -104.5363
18.0127 | -104.5416
18.0164 | -104.5487
18.02 | -104.5553
18.0237 | -104.5626
18.0273 | -104.5685
18.031 | -104.5741
18.0347 | -104.5811
18.0383 | -104.5909
18.042 | -104.601
18.0456 | -104.6067
18.0493 | -104.6136
18.053 | -104.6195
18.0566 | -104.6253
18.0603 | -104.6302
18.0639 | -104.6338
18.0676 | -104.636
18.0712 | -104.638
18.0749 | -104.6402
18.0786 | -104.6417
18.0822 | -104.6443
18.0859 | -104.6478
18.0895 | -104.6464
18.0932 | -104.6447
18.0968 | -104.6426
18.1005 | -104.6384
18.1042 | -104.6307
18.1078 | -104.6233
18.1115 | -104.6169
18.1151 | -104.6078
18.1188 | -104.5971
18.1225 | -104.5883
18.1261 | -104.578
18.1298 | -104.5677
18.1334 | -104.5562
18.1371 | -104.5435
18.1407 | -104.5304
18.1444 | -104.517
18.1481 | -104.5039
18.1517 | -104.4887
18.1554 | -104.4733
18.159 | -104.4585
18.1627 | -104.445
18.1664 | -104.4287
18.17 | -104.4133
18.1737 | -104.4012
18.1773 | -104.3886
18.181 | -104.3776
18.1846 | -104.3675
18.1883 | -104.3576
18.192 | -104.3474
18.1956 | -104.3365
18.1993 | -104.3255
18.2029 | -104.3181
18.2066 | -104.3114
18.2103 | -104.3063
18.2139 | -104.3026
18.2176 | -104.2997
18.2212 | -104.2994
18.2249 | -104.2993
18.2285 | -104.2979
18.2322 | -104.3
18.2359 | -104.2997
18.2395 | -104.2992
18.2432 | -104.2996
18.2468 | -104.302
18.2505 | -104.3053
18.2541 | -104.3095
18.2578 | -104.3128
18.2615 | -104.3188
18.2651 | -104.3242
18.2688 | -104.3268
18.2724 | -104.3305
18.2761 | -104.3355
18.2798 | -104.3413
18.2834 | -104.3457
18.2871 | -104.3492
18.2907 | -104.354
18.2944 | -104.3596
18.298 | -104.3644
18.3017 | -104.3664
18.3054 | -104.3734
18.309 | -104.3783
18.3127 | -104.3818
18.3163 | -104.3845
18.32 | -104.387
18.3237 | -104.3911
18.3273 | -104.3954
18.331 | -104.3996
18.3346 | -104.4036
18.3383 | -104.406
18.3419 | -104.4057
18.3456 | -104.4043
18.3493 | -104.4045
18.3529 | -104.4049
18.3566 | -104.4055
18.3602 | -104.4098
18.3639 | -104.4151
18.3676 | -104.418
18.3712 | -104.4168
18.3749 | -104.4173
18.3785 | -104.4209
18.3822 | -104.4239
18.3858 | -104.4255
18.3895 | -104.4268
18.3932 | -104.4288
18.3968 | -104.4286
18.4005 | -104.4304
18.4041 | -104.4321
18.4078 | -104.4327
18.4114 | -104.4354
18.4151 | -104.4378
18.4188 | -104.4366
18.4224 | -104.4349
18.4261 | -104.4358
18.4297 | -104.4341
18.4334 | -104.43
18.4371 | -104.4278
18.4407 | -104.4235
18.4444 | -104.4186
18.448 | -104.4152
18.4517 | -104.4105
18.4553 | -104.4066
18.459 | -104.4016
18.4627 | -104.3934
18.4663 | -104.3853
18.47 | -104.3758
18.4736 | -104.3662
18.4773 | -104.3563
18.481 | -104.3464
18.4846 | -104.3349
18.4883 | -104.3233
18.4919 | -104.3128
18.4956 | -104.3002
18.4992 | -104.292
18.5029 | -104.2837
18.5066 | -104.2754
18.5102 | -104.2676
18.5139 | -104.258
18.5175 | -104.2458
18.5212 | -104.2328
18.5249 | -104.2207
18.5285 | -104.2078
18.5322 | -104.1973
18.5358 | -104.1871
18.5395 | -104.1765
18.5431 | -104.1655
18.5468 | -104.1577
18.5505 | -104.1506
18.5541 | -104.1406
18.5578 | -104.13
18.5614 | -104.1222
18.5651 | -104.116
18.5688 | -104.1062
18.5724 | -104.0965
18.5761 | -104.087
18.5797 | -104.078
18.5834 | -104.0709
18.587 | -104.0637
18.5907 | -104.0579
18.5944 | -104.0499
18.598 | -104.041
18.6017 | -104.0347
18.6053 | -104.0314
18.609 | -104.0276
18.6126 | -104.0217
18.6163 | -104.0162
18.62 | -104.0121
18.6236 | -104.0071
18.6273 | -104.0023
18.6309 | -103.9966
18.6346 | -103.9944
18.6383 | -103.9931
18.6419 | -103.9908
18.6456 | -103.9883
18.6492 | -103.9875
18.6529 | -103.9861
18.6565 | -103.9832
18.6602 | -103.9828
18.6639 | -103.9828
18.6675 | -103.9825
18.6712 | -103.9824
18.6748 | -103.978
18.6785 | -103.9749
18.6822 | -103.9737
18.6858 | -103.9698
18.6895 | -103.9667
18.6931 | -103.964
18.6968 | -103.9595
18.7004 | -103.954
18.7041 | -103.9483
18.7078 | -103.9451
18.7114 | -103.9386
18.7151 | -103.9309
18.7187 | -103.9237
18.7224 | -103.9162
18.7261 | -103.9081
18.7297 | -103.8977
18.7334 | -103.8872
18.737 | -103.877
18.7407 | -103.8684
18.7443 | -103.858
18.748 | -103.8468
18.7517 | -103.838
18.7553 | -103.8296
18.759 | -103.8193
18.7626 | -103.8089
18.7663 | -103.7958
18.7699 | -103.7837
18.7736 | -103.7709
18.7773 | -103.7585
18.7809 | -103.7493
18.7846 | -103.7417
18.7882 | -103.7344
18.7919 | -103.7238
18.7956 | -103.7151
18.7992 | -103.7074
18.8029 | -103.6993
18.8065 | -103.6904
18.8102 | -103.6832
18.8138 | -103.678
18.8175 | -103.6734
18.8212 | -103.668
18.8248 | -103.6617
18.8285 | -103.6548
18.8321 | -103.6492
18.8358 | -103.6464
18.8395 | -103.6455
18.8431 | -103.6461
18.8468 | -103.6461
18.8504 | -103.6443
18.8541 | -103.6423
18.8577 | -103.6401
18.8614 | -103.6418
18.8651 | -103.6427
18.8687 | -103.6411
18.8724 | -103.64
18.876 | -103.6374
18.8797 | -103.6358
18.8834 | -103.6366
18.887 | -103.6372
18.8907 | -103.6365
18.8943 | -103.6355
18.898 | -103.6374
18.9016 | -103.6393
18.9053 | -103.639
18.909 | -103.6386
18.9126 | -103.6376
18.9163 | -103.6343
18.9199 | -103.6306
18.9236 | -103.6308
18.9272 | -103.6324
18.9309 | -103.6304
18.9346 | -103.63
18.9382 | -103.6298
18.9419 | -103.6307
18.9455 | -103.6313
18.9492 | -103.6313
18.9529 | -103.6307
18.9565 | -103.6269
18.9602 | -103.6228
18.9638 | -103.62
18.9675 | -103.6169
18.9711 | -103.6143
18.9748 | -103.6101
18.9785 | -103.6079
18.9821 | -103.6048
18.9858 | -103.6021
18.9894 | -103.5984
18.9931 | -103.596
18.9968 | -103.59
19.0004 | -103.5875
19.0041 | -103.5838
19.0077 | -103.581
19.0114 | -103.5762
19.015 | -103.5703
19.0187 | -103.5669
19.0224 | -103.5638
19.026 | -103.5607
19.0297 | -103.5581
19.0333 | -103.5556
19.037 | -103.554
19.0407 | -103.5505
19.0443 | -103.5475
19.048 | -103.5459
19.0516 | -103.5443
19.0553 | -103.5434
19.0589 | -103.5427
19.0626 | -103.5451
19.0663 | -103.5471
19.0699 | -103.5477
19.0736 | -103.5488
19.0772 | -103.552
19.0809 | -103.5519
19.0845 | -103.55
19.0882 | -103.5496
19.0919 | -103.5518
19.0955 | -103.5544
19.0992 | -103.5577
19.1028 | -103.5607
19.1065 | -103.5626
19.1102 | -103.5642
19.1138 | -103.5664
19.1175 | -103.5686
19.1211 | -103.569
19.1248 | -103.5704
19.1284 | -103.572
19.1321 | -103.5702
19.1358 | -103.5697
19.1394 | -103.5705
19.1431 | -103.5696
19.1467 | -103.5683
19.1504 | -103.5652
19.1541 | -103.5601
19.1577 | -103.5567
19.1614 | -103.5527
19.165 | -103.5479
19.1687 | -103.5425
19.1723 | -103.5366
19.176 | -103.5312
19.1797 | -103.525
19.1833 | -103.5192
19.187 | -103.5134
19.1906 | -103.5075
19.1943 | -103.4982
19.198 | -103.4877
19.2016 | -103.4786
19.2053 | -103.4682
19.2089 | -103.4581
19.2126 | -103.4504
19.2162 | -103.4401
19.2199 | -103.4294
19.2236 | -103.421
19.2272 | -103.4123
19.2309 | -103.4031
19.2345 | -103.3946
19.2382 | -103.3854
19.2418 | -103.3764
19.2455 | -103.367
19.2492 | -103.3587
19.2528 | -103.3517
19.2565 | -103.3454
19.2601 | -103.3392
19.2638 | -103.331
19.2675 | -103.3259
19.2711 | -103.3236
19.2748 | -103.3232
19.2784 | -103.3229
19.2821 | -103.3193
19.2857 | -103.3169
19.2894 | -103.317
19.2931 | -103.3171
19.2967 | -103.3149
19.3004 | -103.3145
19.304 | -103.3155
19.3077 | -103.3146
19.3114 | -103.3126
19.315 | -103.3142
19.3187 | -103.3155
19.3223 | -103.3157
19.326 | -103.3165
19.3296 | -103.3148
19.3333 | -103.3131
19.337 | -103.3113
19.3406 | -103.3086
19.3443 | -103.3067
19.3479 | -103.3056
19.3516 | -103.3036
19.3553 | -103.3036
19.3589 | -103.3018
19.3626 | -103.3
19.3662 | -103.2978
19.3699 | -103.2937
19.3735 | -103.2899
19.3772 | -103.2868
19.3809 | -103.2838
19.3845 | -103.2801
19.3882 | -103.276
19.3918 | -103.2729
19.3955 | -103.2729
19.3991 | -103.272
19.4028 | -103.2714
19.4065 | -103.2716
19.4101 | -103.2702
19.4138 | -103.2688
19.4174 | -103.2685
19.4211 | -103.2688
19.4248 | -103.2702
19.4284 | -103.2699
19.4321 | -103.2695
19.4357 | -103.2711
19.4394 | -103.2727
19.443 | -103.2742
19.4467 | -103.2746
19.4504 | -103.276
19.454 | -103.2779
19.4577 | -103.2812
19.4613 | -103.2851
19.465 | -103.2879
19.4687 | -103.2917
19.4723 | -103.2938
19.476 | -103.2975
19.4796 | -103.3012
19.4833 | -103.3052
19.4869 | -103.3081
19.4906 | -103.3114
19.4943 | -103.3142
19.4979 | -103.3152
19.5016 | -103.3165
19.5052 | -103.3181
19.5089 | -103.3173
19.5126 | -103.3159
19.5162 | -103.3153
19.5199 | -103.3138
19.5235 | -103.3078
19.5272 | -103.3078
19.5308 | -103.3082
19.5345 | -103.3044
19.5382 | -103.3002
19.5418 | -103.2985
19.5455 | -103.2957
19.5491 | -103.2918
19.5528 | -103.2868
19.5565 | -103.2812
19.5601 | -103.278
19.5638 | -103.2745
19.5674 | -103.2699
19.5711 | -103.2653
19.5747 | -103.2631
19.5784 | -103.2597
19.5821 | -103.2555
19.5857 | -103.2538
19.5894 | -103.2504
19.593 | -103.2438
19.5967 | -103.2376
19.6003 | -103.2302
19.604 | -103.2217
19.6077 | -103.2165
19.6113 | -103.2113
19.615 | -103.2031
19.6186 | -103.1945
19.6223 | -103.1861
19.626 | -103.1783
19.6296 | -103.1697
19.6333 | -103.1614
19.6369 | -103.1544
19.6406 | -103.1461
19.6442 | -103.1389
19.6479 | -103.1332
19.6516 | -103.1249
19.6552 | -103.1157
19.6589 | -103.1089
19.6625 | -103.1017
19.6662 | -103.0932
19.6699 | -103.0866
19.6735 | -103.0796
19.6772 | -103.0728
19.6808 | -103.0681
19.6845 | -103.0652
19.6881 | -103.0624
19.6918 | -103.0599
19.6955 | -103.0584
19.6991 | -103.0588
19.7028 | -103.0582
19.7064 | -103.0587
19.7101 | -103.0595
19.7138 | -103.0599
19.7174 | -103.0618
19.7211 | -103.0643
19.7247 | -103.0687
19.7284 | -103.0727
19.732 | -103.0754
19.7357 | -103.0779
19.7394 | -103.0834
19.743 | -103.092
19.7467 | -103.0997
19.7503 | -103.1089
19.754 | -103.1191
19.7576 | -103.1301
19.7613 | -103.1412
19.765 | -103.1529
19.7686 | -103.1661
19.7723 | -103.1806
19.7759 | -103.1955
19.7796 | -103.2125
19.7833 | -103.229
19.7869 | -103.2462
19.7906 | -103.2626
19.7942 | -103.2791
19.7979 | -103.2979
19.8015 | -103.318
19.8052 | -103.3356
19.8089 | -103.3505
19.8125 | -103.3639
19.8162 | -103.3804
19.8198 | -103.3974
19.8235 | -103.414
19.8272 | -103.4294
19.8308 | -103.4414
19.8345 | -103.4545
19.8381 | -103.4674
19.8418 | -103.4792
19.8454 | -103.4896
19.8491 | -103.4987
19.8528 | -103.5057
19.8564 | -103.5101
19.8601 | -103.5139
19.8637 | -103.5161
19.8674 | -103.5192
19.8711 | -103.5233
19.8747 | -103.525
19.8784 | -103.5246
19.882 | -103.5251
19.8857 | -103.524
19.8893 | -103.5184
19.893 | -103.5111
19.8967 | -103.5069
19.9003 | -103.5019
19.904 | -103.4938
19.9076 | -103.4821
19.9113 | -103.4691
19.9149 | -103.4565
19.9186 | -103.4439
19.9223 | -103.4314
19.9259 | -103.4166
19.9296 | -103.4011
19.9332 | -103.3878
19.9369 | -103.3769
19.9406 | -103.3623
19.9442 | -103.3488
19.9479 | -103.3377
19.9515 | -103.3265
19.9552 | -103.3139
19.9588 | -103.3002
19.9625 | -103.2859
19.9662 | -103.2712
19.9698 | -103.2599
19.9735 | -103.2495
19.9771 | -103.2383
19.9808 | -103.2271
19.9845 | -103.2186
19.9881 | -103.2109
19.9918 | -103.2045
19.9954 | -103.1992
19.9991 | -103.1938
20.0027 | -103.1895
20.0064 | -103.186
20.0101 | -103.1842
20.0137 | -103.184
20.0174 | -103.1831
20.021 | -103.1827
20.0247 | -103.1801
20.0284 | -103.1787
20.032 | -103.1782
20.0357 | -103.1772
20.0393 | -103.1776
20.043 | -103.1797
20.0466 | -103.1828
20.0503 | -103.1856
20.054 | -103.1924
20.0576 | -103.1956
20.0613 | -103.1977
20.0649 | -103.2001
20.0686 | -103.2028
20.0722 | -103.2049
20.0759 | -103.2093
20.0796 | -103.2131
20.0832 | -103.2166
20.0869 | -103.2183
20.0905 | -103.2204
20.0942 | -103.224
20.0979 | -103.2286
20.1015 | -103.2326
20.1052 | -103.2359
20.1088 | -103.2366
20.1125 | -103.2377
20.1161 | -103.2388
20.1198 | -103.2367
20.1235 | -103.2374
20.1271 | -103.2378
20.1308 | -103.2399
20.1344 | -103.2401
20.1381 | -103.2391
20.1418 | -103.2402
20.1454 | -103.2386
20.1491 | -103.2374
20.1527 | -103.2346
20.1564 | -103.2321
20.16 | -103.2309
20.1637 | -103.228
20.1674 | -103.2233
20.171 | -103.2209
20.1747 | -103.2195
20.1783 | -103.2181
20.182 | -103.2131
20.1857 | -103.2085
20.1893 | -103.2043
20.193 | -103.1993
20.1966 | -103.1963
20.2003 | -103.1924
20.2039 | -103.1867
20.2076 | -103.1814
20.2113 | -103.1779
20.2149 | -103.1744
20.2186 | -103.1684
20.2222 | -103.1603
20.2259 | -103.1545
20.2295 | -103.15
20.2332 | -103.147
20.2369 | -103.1418
20.2405 | -103.1382
20.2442 | -103.1368
20.2478 | -103.1344
20.2515 | -103.1336
20.2552 | -103.1306
20.2588 | -103.1261
20.2625 | -103.1232
20.2661 | -103.1218
20.2698 | -103.1201
20.2734 | -103.1183
20.2771 | -103.1177
20.2808 | -103.1176
20.2844 | -103.1171
20.2881 | -103.1171
20.2917 | -103.1182
20.2954 | -103.1197
20.2991 | -103.1228
20.3027 | -103.1271
20.3064 | -103.1295
20.31 | -103.1347
20.3137 | -103.1417
20.3173 | -103.1482
20.321 | -103.1553
20.3247 | -103.1632
20.3283 | -103.1706
20.332 | -103.1801
20.3356 | -103.1921
20.3393 | -103.2016
20.343 | -103.2122
20.3466 | -103.2246
20.3503 | -103.2384
20.3539 | -103.2536
20.3576 | -103.2697
20.3612 | -103.2869
20.3649 | -103.3047
20.3686 | -103.3232
20.3722 | -103.3395
20.3759 | -103.3562
20.3795 | -103.3751
20.3832 | -103.3954
20.3868 | -103.4142
20.3905 | -103.4356
20.3942 | -103.4551
20.3978 | -103.4746
20.4015 | -103.4938
20.4051 | -103.5131
20.4088 | -103.5332
20.4125 | -103.5489
20.4161 | -103.566
20.4198 | -103.5819
20.4234 | -103.5964
20.4271 | -103.6101
20.4307 | -103.6191
20.4344 | -103.6326
20.4381 | -103.6458
20.4417 | -103.6557
20.4454 | -103.6619
20.449 | -103.6661
20.4527 | -103.6716
20.4564 | -103.6768
20.46 | -103.6761
20.4637 | -103.6752
20.4673 | -103.6738
20.471 | -103.6742
20.4746 | -103.673
20.4783 | -103.6693
20.482 | -103.665
20.4856 | -103.6585
20.4893 | -103.6489
20.4929 | -103.6392
20.4966 | -103.6312
20.5003 | -103.6195
20.5039 | -103.6081
20.5076 | -103.5984
20.5112 | -103.5873
20.5149 | -103.5761
20.5185 | -103.5628
20.5222 | -103.5498
20.5259 | -103.54
20.5295 | -103.5305
20.5332 | -103.5191
20.5368 | -103.5083
20.5405 | -103.4982
20.5441 | -103.4908
20.5478 | -103.4861
20.5515 | -103.4804
20.5551 | -103.4739
20.5588 | -103.4675
20.5624 | -103.4622
20.5661 | -103.456
20.5698 | -103.4513
20.5734 | -103.4478
20.5771 | -103.4458
20.5807 | -103.4519
20.5844 | -103.4514
20.588 | -103.4495
20.5917 | -103.4489
20.5954 | -103.4489
20.599 | -103.4491
20.6027 | -103.449
20.6063 | -103.4492
20.61 | -103.4505
20.6137 | -103.4534
20.6173 | -103.4566
20.621 | -103.4611
20.6246 | -103.4634
20.6283 | -103.4666
20.6319 | -103.474
20.6356 | -103.48
20.6393 | -103.4867
20.6429 | -103.4955
20.6466 | -103.502
20.6502 | -103.5072
20.6539 | -103.5131
20.6576 | -103.5197
20.6612 | -103.5238
20.6649 | -103.5302
20.6685 | -103.5385
20.6722 | -103.5461
20.6758 | -103.5515
20.6795 | -103.5595
20.6832 | -103.5663
20.6868 | -103.5754
20.6905 | -103.5838
20.6941 | -103.5919
20.6978 | -103.6006
20.7015 | -103.6078
20.7051 | -103.6169
20.7088 | -103.6253
20.7124 | -103.6323
20.7161 | -103.6414
20.7197 | -103.6492
20.7234 | -103.6587
20.7271 | -103.665
20.7307 | -103.6712
20.7344 | -103.6791
20.738 | -103.6883
20.7417 | -103.6967
20.7453 | -103.7035
20.749 | -103.7111
20.7527 | -103.722
20.7563 | -103.733
20.76 | -103.7428
20.7636 | -103.7515
20.7673 | -103.7596
20.771 | -103.7691
20.7746 | -103.7792
20.7783 | -103.7892
20.7819 | -103.803
20.7856 | -103.8158
20.7892 | -103.8263
20.7929 | -103.8385
20.7966 | -103.8482
20.8002 | -103.8567
20.8039 | -103.8691
20.8075 | -103.881
20.8112 | -103.8921
20.8149 | -103.9012
20.8185 | -103.909
20.8222 | -103.9178
20.8258 | -103.9275
20.8295 | -103.9386
20.8331 | -103.9472
20.8368 | -103.9551
20.8405 | -103.9624
20.8441 | -103.9712
20.8478 | -103.9809
20.8514 | -103.9891
20.8551 | -103.9983
20.8588 | -104.0077
20.8624 | -104.0152
20.8661 | -104.025
20.8697 | -104.0332
20.8734 | -104.0394
20.877 | -104.0453
20.8807 | -104.0513
20.8844 | -104.0603
20.888 | -104.0677
20.8917 | -104.0741
20.8953 | -104.0817
20.899 | -104.09
20.9026 | -104.0967
20.9063 | -104.1015
20.91 | -104.1046
20.9136 | -104.1088
20.9173 | -104.1134
20.9209 | -104.1199
20.9246 | -104.1255
20.9283 | -104.1288
20.9319 | -104.1322
20.9356 | -104.1357
20.9392 | -104.1378
20.9429 | -104.1405
20.9465 | -104.1473
20.9502 | -104.1553
20.9539 | -104.1603
20.9575 | -104.1661
20.9612 | -104.173
20.9648 | -104.1788
20.9685 | -104.1837
20.9722 | -104.1869
20.9758 | -104.1893
20.9795 | -104.1933
20.9831 | -104.1969
20.9868 | -104.2011
20.9904 | -104.2054
20.9941 | -104.209
20.9978 | -104.2132
21.0014 | -104.2167
21.0051 | -104.22
21.0087 | -104.2253
21.0124 | -104.2301
21.016 | -104.2355
21.0197 | -104.241
21.0234 | -104.2459
21.027 | -104.2496
21.0307 | -104.2538
21.0343 | -104.2582
21.038 | -104.2655
21.0417 | -104.2747
21.0453 | -104.2829
21.049 | -104.2902
21.0526 | -104.2981
21.0563 | -104.3079
21.0599 | -104.3185
21.0636 | -104.3301
21.0673 | -104.3423
21.0709 | -104.3556
21.0746 | -104.3699
21.0782 | -104.3825
21.0819 | -104.3957
21.0856 | -104.4097
21.0892 | -104.4243
21.0929 | -104.4435
21.0965 | -104.4637
21.1002 | -104.4813
21.1038 | -104.498
21.1075 | -104.5179
21.1112 | -104.5374
21.1148 | -104.5566
21.1185 | -104.5768
21.1221 | -104.5959
21.1258 | -104.6167
21.1295 | -104.6354
21.1331 | -104.6546
21.1368 | -104.6735
21.1404 | -104.692
21.1441 | -104.7099
21.1477 | -104.7276
21.1514 | -104.7495
21.1551 | -104.7676
21.1587 | -104.785
21.1624 | -104.804
21.166 | -104.8198
21.1697 | -104.8364
21.1733 | -104.8536
21.177 | -104.8708
21.1807 | -104.8878
21.1843 | -104.9028
21.188 | -104.9182
21.1916 | -104.9354
21.1953 | -104.951
21.199 | -104.9648
21.2026 | -104.9792
21.2063 | -104.9937
21.2099 | -105.0079
21.2136 | -105.0208
21.2172 | -105.0325
21.2209 | -105.0432
21.2246 | -105.0528
21.2282 | -105.0617
21.2319 | -105.0724
21.2355 | -105.0834
21.2392 | -105.0923
21.2429 | -105.097
21.2465 | -105.102
21.2502 | -105.1051
21.2538 | -105.1093
21.2575 | -105.1132
21.2611 | -105.1141
21.2648 | -105.1159
21.2685 | -105.117
21.2721 | -105.1197
21.2758 | -105.1219
21.2794 | -105.1194
21.2831 | -105.116
21.2868 | -105.113
21.2904 | -105.1095
21.2941 | -105.1063
21.2977 | -105.103
21.3014 | -105.0983
21.305 | -105.0912
21.3087 | -105.0834
21.3124 | -105.078
21.316 | -105.0721
21.3197 | -105.0644
21.3233 | -105.0581
21.327 | -105.0505
21.3307 | -105.0423
21.3343 | -105.0351
21.338 | -105.0257
21.3416 | -105.013
21.3453 | -105.0018
21.3489 | -104.9931
21.3526 | -104.9844
21.3563 | -104.9764
21.3599 | -104.9695
21.3636 | -104.9619
21.3672 | -104.9552
21.3709 | -104.9489
21.3745 | -104.9424
21.3782 | -104.9353
21.3819 | -104.9308
21.3855 | -104.9269
21.3892 | -104.9248
21.3928 | -104.9201
21.3965 | -104.9155
21.4002 | -104.9113
21.4038 | -104.9083
21.4075 | -104.905
21.4111 | -104.9021
21.4148 | -104.9013
21.4184 | -104.9023
21.4221 | -104.9042
21.4258 | -104.9071
21.4294 | -104.9098
21.4331 | -104.9138
21.4367 | -104.9183
21.4404 | -104.9236
21.4441 | -104.9308
21.4477 | -104.9371
21.4514 | -104.9429
21.455 | -104.9477
21.4587 | -104.9567
21.4623 | -104.9659
21.466 | -104.9741
21.4697 | -104.9853
21.4733 | -104.9951
21.477 | -105.0025
21.4806 | -105.0117
21.4843 | -105.0207
21.488 | -105.0296
21.4916 | -105.0386
21.4953 | -105.0465
21.4989 | -105.0553
21.5026 | -105.0647
21.5062 | -105.0746
21.5099 | -105.0846
21.5136 | -105.0942
21.5172 | -105.1014
21.5209 | -105.1087
21.5245 | -105.1165
21.5282 | -105.1235
21.5318 | -105.1308
21.5355 | -105.1374
21.5392 | -105.1411
21.5428 | -105.1455
21.5465 | -105.1545
21.5501 | -105.1602
21.5538 | -105.1644
21.5575 | -105.1677
21.5611 | -105.1733
21.5648 | -105.1795
21.5684 | -105.1828
21.5721 | -105.1868
21.5757 | -105.191
21.5794 | -105.1955
21.5831 | -105.2005
21.5867 | -105.2067
21.5904 | -105.2135
21.594 | -105.2179
21.5977 | -105.2217
21.6014 | -105.2247
21.605 | -105.227
21.6087 | -105.2333
21.6123 | -105.2392
21.616 | -105.2409
21.6196 | -105.2429
21.6233 | -105.2467
21.627 | -105.2488
21.6306 | -105.2521
21.6343 | -105.2561
21.6379 | -105.263
21.6416 | -105.268
21.6453 | -105.2727
21.6489 | -105.2784
21.6526 | -105.2806
21.6562 | -105.2833
21.6599 | -105.2868
21.6635 | -105.2897
21.6672 | -105.2936
21.6709 | -105.2981
21.6745 | -105.303
21.6782 | -105.3059
21.6818 | -105.3091
21.6855 | -105.3127
21.6891 | -105.315
21.6928 | -105.3183
21.6965 | -105.322
21.7001 | -105.326
21.7038 | -105.3289
21.7074 | -105.3341
21.7111 | -105.3392
21.7148 | -105.3443
21.7184 | -105.3473
21.7221 | -105.3482
21.7257 | -105.348
21.7294 | -105.3471
21.733 | -105.3459
21.7367 | -105.3458
21.7404 | -105.3462
21.744 | -105.3469
21.7477 | -105.3457
21.7513 | -105.3433
21.755 | -105.3427
21.7587 | -105.3424
21.7623 | -105.3415
21.766 | -105.3383
21.7696 | -105.3344
21.7733 | -105.3325
21.7769 | -105.328
21.7806 | -105.3221
21.7843 | -105.3191
21.7879 | -105.3145
21.7916 | -105.3116
21.7952 | -105.3081
21.7989 | -105.3012
21.8026 | -105.2934
21.8062 | -105.2853
21.8099 | -105.2777
21.8135 | -105.2695
21.8172 | -105.2624
21.8208 | -105.2563
21.8245 | -105.2512
21.8282 | -105.2446
21.8318 | -105.2367
21.8355 | -105.2292
21.8391 | -105.221
21.8428 | -105.2136
21.8464 | -105.2083
21.8501 | -105.2026
21.8538 | -105.1958
21.8574 | -105.1888
21.8611 | -105.1828
21.8647 | -105.177
21.8684 | -105.1707
21.8721 | -105.1663
21.8757 | -105.1622
21.8794 | -105.1573
21.883 | -105.1525
21.8867 | -105.1488
21.8903 | -105.1456
21.894 | -105.1418
21.8977 | -105.1368
21.9013 | -105.1344
21.905 | -105.1338
21.9086 | -105.1343
21.9123 | -105.1359
21.916 | -105.1327
21.9196 | -105.128
21.9233 | -105.1288
21.9269 | -105.1287
21.9306 | -105.1276
21.9342 | -105.1276
21.9379 | -105.1285
21.9416 | -105.1265
21.9452 | -105.1243
21.9489 | -105.1236
21.9525 | -105.1221
21.9562 | -105.1209
21.9599 | -105.1203
21.9635 | -105.1207
21.9672 | -105.1193
21.9708 | -105.1184
21.9745 | -105.1186
21.9781 | -105.117
21.9818 | -105.1157
21.9855 | -105.114
21.9891 | -105.1147
21.9928 | -105.115
21.9964 | -105.1158
22.0001 | -105.1179
22.0038 | -105.117
22.0074 | -105.1143
22.0111 | -105.1133
22.0147 | -105.1127
22.0184 | -105.1102
22.022 | -105.1098
22.0257 | -105.1101
22.0294 | -105.1108
22.033 | -105.1105
22.0367 | -105.1113
22.0403 | -105.1119
22.044 | -105.1119
22.0476 | -105.1131
22.0513 | -105.113
22.055 | -105.1103
22.0586 | -105.1096
22.0623 | -105.1098
22.0659 | -105.108
22.0696 | -105.1051
22.0733 | -105.1034
22.0769 | -105.1024
22.0806 | -105.1034
22.0842 | -105.1013
22.0879 | -105.1015
22.0915 | -105.1022
22.0952 | -105.1009
22.0989 | -105.1017
22.1025 | -105.102
22.1062 | -105.1005
22.1098 | -105.0997
22.1135 | -105.0988
22.1172 | -105.0982
22.1208 | -105.0952
22.1245 | -105.0916
22.1281 | -105.0875
22.1318 | -105.0832
22.1354 | -105.0784
22.1391 | -105.0759
22.1428 | -105.0738
22.1464 | -105.0718
22.1501 | -105.0717
22.1537 | -105.0691
22.1574 | -105.067
22.161 | -105.0652
22.1647 | -105.0652
22.1684 | -105.0647
22.172 | -105.0651
22.1757 | -105.0649
22.1793 | -105.0667
22.183 | -105.0695
22.1867 | -105.0725
22.1903 | -105.076
22.194 | -105.0803
22.1976 | -105.0845
22.2013 | -105.089
22.2049 | -105.0948
22.2086 | -105.1003
22.2123 | -105.1058
22.2159 | -105.113
22.2196 | -105.1192
22.2232 | -105.1252
22.2269 | -105.1335
22.2306 | -105.1437
22.2342 | -105.1552
22.2379 | -105.1649
22.2415 | -105.1775
22.2452 | -105.1905
22.2488 | -105.2032
22.2525 | -105.2152
22.2562 | -105.2285
22.2598 | -105.2396
22.2635 | -105.2503
22.2671 | -105.2609
22.2708 | -105.2731
22.2745 | -105.2844
22.2781 | -105.2962
22.2818 | -105.3086
22.2854 | -105.3182
22.2891 | -105.3277
22.2927 | -105.3375
22.2964 | -105.3478
22.3001 | -105.3581
22.3037 | -105.3653
22.3074 | -105.3727
22.311 | -105.3787
22.3147 | -105.3838
22.3183 | -105.39
22.322 | -105.3961
22.3257 | -105.4019
22.3293 | -105.4077
22.333 | -105.4098
22.3366 | -105.4088
22.3403 | -105.4083
22.344 | -105.4086
22.3476 | -105.4072
22.3513 | -105.4044
22.3549 | -105.4035
22.3586 | -105.4043
22.3622 | -105.4021
22.3659 | -105.3992
22.3696 | -105.3975
22.3732 | -105.3941
22.3769 | -105.3902
22.3805 | -105.3866
22.3842 | -105.3846
22.3879 | -105.3808
22.3915 | -105.3779
22.3952 | -105.3743
22.3988 | -105.3703
22.4025 | -105.3665
22.4061 | -105.3655
22.4098 | -105.3644
22.4135 | -105.3617
22.4171 | -105.3592
22.4208 | -105.3597
22.4244 | -105.3618
22.4281 | -105.3639
22.4318 | -105.365
22.4354 | -105.3678
22.4391 | -105.3711
22.4427 | -105.3734
22.4464 | -105.3739
22.45 | -105.3732
22.4537 | -105.3645
22.4574 | -105.3542
22.461 | -105.3388
22.4647 | -105.3157
22.4683 | -105.2966
22.472 | -105.282
22.4757 | -105.2723
22.4793 | -105.2621
22.483 | -105.2496
22.4866 | -105.2358
22.4903 | -105.2198
22.4939 | -105.2027
22.4976 | -105.183
22.5013 | -105.1608
22.5049 | -105.1375
22.5086 | -105.1147
22.5122 | -105.0896
22.5159 | -105.0633
22.5195 | -105.0355
22.5232 | -105.0069
22.5269 | -104.9782
22.5305 | -104.9472
22.5342 | -104.9125
22.5378 | -104.8779
22.5415 | -104.8433
22.5452 | -104.8061
22.5488 | -104.7685
22.5525 | -104.7315
22.5561 | -104.6946
22.5598 | -104.6567
22.5634 | -104.6193
22.5671 | -104.5822
22.5708 | -104.5471
22.5744 | -104.5145
22.5781 | -104.4849
22.5817 | -104.4562
22.5854 | -104.4301
22.5891 | -104.4057
22.5927 | -104.3824
22.5964 | -104.3614
22.6 | -104.3438
22.6037 | -104.3295
22.6073 | -104.3184
22.611 | -104.311
22.6147 | -104.3064
22.6183 | -104.3036
22.622 | -104.3049
22.6256 | -104.3098
22.6293 | -104.3189
22.633 | -104.3298
22.6366 | -104.3429
22.6403 | -104.3595
22.6439 | -104.3777
22.6476 | -104.3988
22.6512 | -104.4247
22.6549 | -104.4532
22.6586 | -104.4817
22.6622 | -104.5128
22.6659 | -104.5463
22.6695 | -104.5809
22.6732 | -104.6157
22.6768 | -104.6516
22.6805 | -104.6887
22.6842 | -104.7276
22.6878 | -104.7657
22.6915 | -104.8004
22.6951 | -104.8391
22.6988 | -104.8772
22.7025 | -104.9139
22.7061 | -104.9495
22.7098 | -104.9839
22.7134 | -105.0167
22.7171 | -105.0485
22.7207 | -105.0792
22.7244 | -105.1073
22.7281 | -105.1346
22.7317 | -105.1617
22.7354 | -105.1876
22.739 | -105.2117
22.7427 | -105.234
22.7464 | -105.255
22.75 | -105.2742
22.7537 | -105.2932
22.7573 | -105.3105
22.761 | -105.324
22.7646 | -105.3381
22.7683 | -105.3519
22.772 | -105.3666
22.7756 | -105.3865
22.7793 | -105.4051
22.7829 | -105.4257
22.7866 | -105.4475
22.7903 | -105.4597
22.7939 | -105.4631
22.7976 | -105.462
22.8012 | -105.4608
22.8049 | -105.4602
22.8085 | -105.4583
22.8122 | -105.4579
22.8159 | -105.4585
22.8195 | -105.4573
22.8232 | -105.457
22.8268 | -105.4572
22.8305 | -105.4562
22.8341 | -105.4555
22.8378 | -105.4539
22.8415 | -105.4532
22.8451 | -105.451
22.8488 | -105.4497
22.8524 | -105.4517
22.8561 | -105.454
22.8598 | -105.455
22.8634 | -105.4549
22.8671 | -105.4556
22.8707 | -105.4583
22.8744 | -105.4607
22.878 | -105.462
22.8817 | -105.4633
22.8854 | -105.4658
22.889 | -105.469
22.8927 | -105.47
22.8963 | -105.4717
22.9 | -105.4721
22.9037 | -105.4737
22.9073 | -105.4756
22.911 | -105.4763
22.9146 | -105.4783
22.9183 | -105.4793
22.9219 | -105.4792
22.9256 | -105.4799
22.9293 | -105.4795
22.9329 | -105.4789
22.9366 | -105.4793
22.9402 | -105.4797
22.9439 | -105.4789
22.9476 | -105.4777
22.9512 | -105.4759
22.9549 | -105.4744
22.9585 | -105.4737
22.9622 | -105.4733
22.9658 | -105.4742
22.9695 | -105.473
22.9732 | -105.4713
22.9768 | -105.4708
22.9805 | -105.4714
22.9841 | -105.4702
22.9878 | -105.469
22.9914 | -105.4689
22.9951 | -105.4692
22.9988 | -105.4678
23.0024 | -105.4674
23.0061 | -105.4679
23.0097 | -105.4695
23.0134 | -105.4709
23.0171 | -105.4711
23.0207 | -105.4703
23.0244 | -105.4692
23.028 | -105.4707
23.0317 | -105.4726
23.0353 | -105.475
23.039 | -105.4783
23.0427 | -105.483
23.0463 | -105.4893
23.05 | -105.495
23.0536 | -105.4985
23.0573 | -105.5016
23.061 | -105.5063
23.0646 | -105.5109
23.0683 | -105.5153
23.0719 | -105.5195
23.0756 | -105.5259
23.0792 | -105.531
23.0829 | -105.5349
23.0866 | -105.5407
23.0902 | -105.5444
23.0939 | -105.5487
23.0975 | -105.5533
23.1012 | -105.558
23.1049 | -105.5621
23.1085 | -105.5662
23.1122 | -105.5714
23.1158 | -105.5785
23.1195 | -105.5839
23.1231 | -105.5864
23.1268 | -105.5888
23.1305 | -105.5926
23.1341 | -105.5965
23.1378 | -105.5982
23.1414 | -105.6004
23.1451 | -105.6047
23.1487 | -105.6069
23.1524 | -105.6084
23.1561 | -105.6116
23.1597 | -105.6133
23.1634 | -105.616
23.167 | -105.6179
23.1707 | -105.6227
23.1744 | -105.6267
23.178 | -105.6271
23.1817 | -105.627
23.1853 | -105.6284
23.189 | -105.6315
23.1926 | -105.6335
23.1963 | -105.6343
23.2 | -105.6364
23.2036 | -105.638
23.2073 | -105.6392
23.2109 | -105.6407
23.2146 | -105.6424
23.2183 | -105.6441
23.2219 | -105.6488
23.2256 | -105.6532
23.2292 | -105.6571
23.2329 | -105.6603
23.2365 | -105.6632
23.2402 | -105.6664
23.2439 | -105.67
23.2475 | -105.6722
23.2512 | -105.674
23.2548 | -105.6754
23.2585 | -105.6769
23.2622 | -105.6783
23.2658 | -105.6793
23.2695 | -105.6821
23.2731 | -105.6849
23.2768 | -105.6876
23.2804 | -105.6894
23.2841 | -105.6907
23.2878 | -105.6917
23.2914 | -105.6919
23.2951 | -105.691
23.2987 | -105.6912
23.3024 | -105.6928
23.306 | -105.6934
23.3097 | -105.6942
23.3134 | -105.6941
23.317 | -105.6949
23.3207 | -105.6943
23.3243 | -105.6926
23.328 | -105.6923
23.3317 | -105.6913
23.3353 | -105.6876
23.339 | -105.6861
23.3426 | -105.6846
23.3463 | -105.6823
23.3499 | -105.6816
23.3536 | -105.6823
23.3573 | -105.681
23.3609 | -105.6789
23.3646 | -105.6772
23.3682 | -105.6749
23.3719 | -105.6724
23.3756 | -105.6726
23.3792 | -105.6718
23.3829 | -105.6712
23.3865 | -105.671
23.3902 | -105.6692
23.3938 | -105.6691
23.3975 | -105.6707
23.4012 | -105.6734
23.4048 | -105.6756
23.4085 | -105.6779
23.4121 | -105.6815
23.4158 | -105.685
23.4195 | -105.6896
23.4231 | -105.6941
23.4268 | -105.6976
23.4304 | -105.7019
23.4341 | -105.7079
23.4377 | -105.7141
23.4414 | -105.7208
23.4451 | -105.7276
23.4487 | -105.7365
23.4524 | -105.7464
23.456 | -105.7556
23.4597 | -105.7644
23.4633 | -105.7725
23.467 | -105.7842
23.4707 | -105.7965
23.4743 | -105.8092
23.478 | -105.8225
23.4816 | -105.8347
23.4853 | -105.8456
23.489 | -105.8584
23.4926 | -105.8716
23.4963 | -105.8833
23.4999 | -105.8951
23.5036 | -105.9052
23.5072 | -105.9174
23.5109 | -105.9312
23.5146 | -105.9434
23.5182 | -105.954
23.5219 | -105.9652
23.5255 | -105.9767
23.5292 | -105.9866
23.5329 | -105.9953
23.5365 | -106.0059
23.5402 | -106.0148
23.5438 | -106.0227
23.5475 | -106.0331
23.5511 | -106.0415
23.5548 | -106.0473
23.5585 | -106.0531
23.5621 | -106.0591
23.5658 | -106.0642
23.5694 | -106.0692
23.5731 | -106.0759
23.5768 | -106.0831
23.5804 | -106.0909
23.5841 | -106.0958
23.5877 | -106.1012
23.5914 | -106.1057
23.595 | -106.1091
23.5987 | -106.1115
23.6024 | -106.1152
23.606 | -106.1187
23.6097 | -106.1209
23.6133 | -106.1235
23.617 | -106.1267
23.6207 | -106.1294
23.6243 | -106.1326
23.628 | -106.136
23.6316 | -106.1399
23.6353 | -106.1404
23.6389 | -106.1412
23.6426 | -106.1427
23.6463 | -106.1457
23.6499 | -106.1481
23.6536 | -106.1502
23.6572 | -106.1528
23.6609 | -106.1559
23.6645 | -106.1594
23.6682 | -106.1651
23.6719 | -106.1676
23.6755 | -106.1718
23.6792 | -106.177
23.6828 | -106.1794
23.6865 | -106.1824
23.6902 | -106.1877
23.6938 | -106.192
23.6975 | -106.1967
23.7011 | -106.2011
23.7048 | -106.205
23.7084 | -106.2076
23.7121 | -106.2141
23.7158 | -106.2213
23.7194 | -106.2284
23.7231 | -106.2352
23.7267 | -106.2421
23.7304 | -106.2488
23.7341 | -106.2546
23.7377 | -106.2614
23.7414 | -106.27
23.745 | -106.2777
23.7487 | -106.2871
23.7523 | -106.2956
23.756 | -106.3028
23.7597 | -106.3104
23.7633 | -106.3185
23.767 | -106.3259
23.7706 | -106.3307
23.7743 | -106.3365
23.778 | -106.3443
23.7816 | -106.351
23.7853 | -106.3591
23.7889 | -106.3665
23.7926 | -106.3726
23.7962 | -106.378
23.7999 | -106.3854
23.8036 | -106.3901
23.8072 | -106.393
23.8109 | -106.3961
23.8145 | -106.3988
23.8182 | -106.402
23.8218 | -106.4051
23.8255 | -106.4074
23.8292 | -106.41
23.8328 | -106.4123
23.8365 | -106.4125
23.8401 | -106.4143
23.8438 | -106.4155
23.8475 | -106.4147
23.8511 | -106.4141
23.8548 | -106.4139
23.8584 | -106.4135
23.8621 | -106.4139
23.8657 | -106.4132
23.8694 | -106.4131
23.8731 | -106.4101
23.8767 | -106.4074
23.8804 | -106.4069
23.884 | -106.4049
23.8877 | -106.4026
23.8914 | -106.4016
23.895 | -106.3992
23.8987 | -106.3983
23.9023 | -106.3983
23.906 | -106.3962
23.9096 | -106.3921
23.9133 | -106.3893
23.917 | -106.3865
23.9206 | -106.3852
23.9243 | -106.3831
23.9279 | -106.3825
23.9316 | -106.3829
23.9353 | -106.3821
23.9389 | -106.3786
23.9426 | -106.3774
23.9462 | -106.3759
23.9499 | -106.3731
23.9535 | -106.3699
23.9572 | -106.3685
23.9609 | -106.3657
23.9645 | -106.3629
23.9682 | -106.3616
23.9718 | -106.3615
23.9755 | -106.3615
23.9791 | -106.3619
23.9828 | -106.3625
23.9865 | -106.3613
23.9901 | -106.3617
23.9938 | -106.3624
23.9974 | -106.3624
24.0011 | -106.3626
24.0048 | -106.3616
24.0084 | -106.3627
24.0121 | -106.3623
24.0157 | -106.3624
24.0194 | -106.365
24.023 | -106.3664
24.0267 | -106.3672
24.0304 | -106.3705
24.034 | -106.373
24.0377 | -106.3748
24.0413 | -106.377
24.045 | -106.38
24.0487 | -106.3802
24.0523 | -106.3812
24.056 | -106.382
24.0596 | -106.3827
24.0633 | -106.3831
24.0669 | -106.385
24.0706 | -106.386
24.0743 | -106.3853
24.0779 | -106.3851
24.0816 | -106.3864
24.0852 | -106.3855
24.0889 | -106.3852
24.0926 | -106.3853
24.0962 | -106.3866
24.0999 | -106.3869
24.1035 | -106.385
24.1072 | -106.3829
24.1108 | -106.3839
24.1145 | -106.3833
24.1182 | -106.3798
24.1218 | -106.3771
24.1255 | -106.3747
24.1291 | -106.3723
24.1328 | -106.3716
24.1364 | -106.3702
24.1401 | -106.3677
24.1438 | -106.3648
24.1474 | -106.363
24.1511 | -106.3598
24.1547 | -106.3556
24.1584 | -106.3538
24.1621 | -106.3525
24.1657 | -106.3511
24.1694 | -106.3486
24.173 | -106.3449
24.1767 | -106.3407
24.1803 | -106.338
24.184 | -106.3359
24.1877 | -106.3335
24.1913 | -106.331
24.195 | -106.328
24.1986 | -106.3256
24.2023 | -106.322
24.206 | -106.3202
24.2096 | -106.3203
24.2133 | -106.319
24.2169 | -106.3181
24.2206 | -106.316
24.2242 | -106.3123
24.2279 | -106.3093
24.2316 | -106.3069
24.2352 | -106.3052
24.2389 | -106.3031
24.2425 | -106.3026
24.2462 | -106.3027
24.2499 | -106.303
24.2535 | -106.3013
24.2572 | -106.2997
24.2608 | -106.2977
24.2645 | -106.2867
24.2681 | -106.2479
24.2718 | -106.1401
24.2755 | -105.6657
24.2791 | -105.2924
24.2828 | -104.9368
24.2864 | -104.694
24.2901 | -104.4837
24.2937 | -104.2981
24.2974 | -104.1171
24.3011 | -103.9326
24.3047 | -103.7432
24.3084 | -103.548
24.312 | -103.3474
24.3157 | -103.1431
24.3194 | -102.9344
24.323 | -102.723
24.3267 | -102.5104
24.3303 | -102.2976
24.334 | -102.0836
24.3376 | -101.869
24.3413 | -101.6557
24.345 | -101.4441
24.3486 | -101.2352
24.3523 | -101.0291
24.3559 | -100.8264
24.3596 | -100.6275
24.3633 | -100.4325
24.3669 | -100.2426
24.3706 | -100.0587
24.3742 | -99.8805
24.3779 | -99.7086
24.3815 | -99.5434
24.3852 | -99.3848
24.3889 | -99.2332
24.3925 | -99.0892
24.3962 | -98.9523
24.3998 | -98.8238
24.4035 | -98.704
24.4072 | -98.5922
24.4108 | -98.4883
24.4145 | -98.3932
24.4181 | -98.3068
24.4218 | -98.2291
24.4254 | -98.1606
24.4291 | -98.1009
24.4328 | -98.0503
24.4364 | -98.0088
24.4401 | -97.9761
24.4437 | -97.9523
24.4474 | -97.9382
24.451 | -97.9333
24.4547 | -97.9374
24.4584 | -97.9508
24.462 | -97.9736
24.4657 | -98.0054
24.4693 | -98.0463
24.473 | -98.096
24.4767 | -98.1548
24.4803 | -98.2227
24.484 | -98.2995
24.4876 | -98.385
24.4913 | -98.4793
24.4949 | -98.5817
24.4986 | -98.6925
24.5023 | -98.8113
24.5059 | -98.938
24.5096 | -99.0727
24.5132 | -99.2151
24.5169 | -99.365
24.5206 | -99.5219
24.5242 | -99.6853
24.5279 | -99.8548
24.5315 | -100.0304
24.5352 | -100.212
24.5388 | -100.3994
24.5425 | -100.5916
24.5462 | -100.788
24.5498 | -100.9885
24.5535 | -101.1917
24.5571 | -101.3979
24.5608 | -101.6057
24.5645 | -101.8136
24.5681 | -102.0205
24.5718 | -102.2143
24.5754 | -102.3942
24.5791 | -102.5563
24.5827 | -102.6942
24.5864 | -102.8252
24.5901 | -102.9485
24.5937 | -103.0732
24.5974 | -103.1982
24.601 | -103.3135
24.6047 | -103.4152
24.6084 | -103.5054
24.612 | -103.5913
24.6157 | -103.6906
24.6193 | -103.8192
24.623 | -103.9598
24.6266 | -104.1302
24.6303 | -104.2504
24.634 | -104.2396
24.6376 | -104.1741
24.6413 | -104.0676
24.6449 | -103.9459
24.6486 | -103.8191
24.6522 | -103.6891
24.6559 | -103.5566
24.6596 | -103.4237
24.6632 | -103.2913
24.6669 | -103.16
24.6705 | -103.0295
24.6742 | -102.9
24.6779 | -102.773
24.6815 | -102.6489
24.6852 | -102.5287
24.6888 | -102.4116
24.6925 | -102.298
24.6961 | -102.1884
24.6998 | -102.0827
24.7035 | -101.9824
24.7071 | -101.8876
24.7108 | -101.7995
24.7144 | -101.7165
24.7181 | -101.6387
24.7218 | -101.5676
24.7254 | -101.5027
24.7291 | -101.4446
24.7327 | -101.3938
24.7364 | -101.3499
24.74 | -101.3129
24.7437 | -101.2838
24.7474 | -101.2611
24.751 | -101.2444
24.7547 | -101.235
24.7583 | -101.2335
24.762 | -101.2393
24.7657 | -101.2523
24.7693 | -101.2723
24.773 | -101.2995
24.7766 | -101.3341
24.7803 | -101.3756
24.7839 | -101.4244
24.7876 | -101.4807
24.7913 | -101.5438
24.7949 | -101.6133
24.7986 | -101.6889
24.8022 | -101.8091
24.8059 | -101.8937
24.8095 | -101.9839
24.8132 | -102.0794
24.8169 | -102.1796
24.8205 | -102.285
24.8242 | -102.3957
24.8278 | -102.5103
24.8315 | -102.6288
24.8352 | -102.7513
24.8388 | -102.8775
24.8425 | -103.0062
24.8461 | -103.1383
24.8498 | -103.2731
24.8534 | -103.4093
24.8571 | -103.5469
24.8608 | -103.6854
24.8644 | -103.8242
24.8681 | -103.963
24.8717 | -104.1019
24.8754 | -104.2385
24.8791 | -104.375
24.8827 | -104.5092
24.8864 | -104.6406
24.89 | -104.7682
24.8937 | -104.8925
24.8973 | -105.0142
24.901 | -105.1321
24.9047 | -105.2464
24.9083 | -105.3569
24.912 | -105.464
24.9156 | -105.5675
24.9193 | -105.6695
24.923 | -105.7898
24.9266 | -105.9207
24.9303 | -106.0628
24.9339 | -106.2215
24.9376 | -106.3517
24.9412 | -106.4565
24.9449 | -106.5183
24.9486 | -106.5463
24.9522 | -106.5642
24.9559 | -106.5816
24.9595 | -106.5992
24.9632 | -106.6154
24.9668 | -106.6329
24.9705 | -106.6461
24.9742 | -106.6558
24.9778 | -106.6648
24.9815 | -106.6709
24.9851 | -106.6787
24.9888 | -106.6886
24.9925 | -106.6977
24.9961 | -106.7027
24.9998 | -106.704
25.0034 | -106.7065
25.0071 | -106.7114
25.0107 | -106.7156
25.0144 | -106.7192
25.0181 | -106.7228
25.0217 | -106.7239
25.0254 | -106.7246
25.029 | -106.725
25.0327 | -106.7258
25.0364 | -106.7273
25.04 | -106.7285
25.0437 | -106.7294
25.0473 | -106.731
25.051 | -106.7343
25.0546 | -106.7367
25.0583 | -106.7375
25.062 | -106.7366
25.0656 | -106.7367
25.0693 | -106.738
25.0729 | -106.7377
25.0766 | -106.7373
25.0803 | -106.7379
25.0839 | -106.7371
25.0876 | -106.7362
25.0912 | -106.7344
25.0949 | -106.7327
25.0985 | -106.7319
25.1022 | -106.7328
25.1059 | -106.7333
25.1095 | -106.7329
25.1132 | -106.7347
25.1168 | -106.7355
25.1205 | -106.7351
25.1241 | -106.733
25.1278 | -106.7319
25.1315 | -106.7317
25.1351 | -106.7292
25.1388 | -106.7282
25.1424 | -106.7271
25.1461 | -106.7257
25.1498 | -106.7238
25.1534 | -106.7208
25.1571 | -106.7184
25.1607 | -106.7155
25.1644 | -106.7136
25.168 | -106.7108
25.1717 | -106.7082
25.1754 | -106.7069
25.179 | -106.7067
25.1827 | -106.7052
25.1863 | -106.7031
25.19 | -106.6997
25.1937 | -106.6956
25.1973 | -106.6935
25.201 | -106.6918
25.2046 | -106.6886
25.2083 | -106.6854
25.2119 | -106.6822
25.2156 | -106.6787
25.2193 | -106.6739
25.2229 | -106.6693
25.2266 | -106.6651
25.2302 | -106.6625
25.2339 | -106.6591
25.2376 | -106.6569
25.2412 | -106.6547
25.2449 | -106.6524
25.2485 | -106.6507
25.2522 | -106.6488
25.2558 | -106.6455
25.2595 | -106.6435
25.2632 | -106.6411
25.2668 | -106.6391
25.2705 | -106.6386
25.2741 | -106.6385
25.2778 | -106.6388
25.2814 | -106.6404
25.2851 | -106.6422
25.2888 | -106.6424
25.2924 | -106.6425
25.2961 | -106.6444
25.2997 | -106.6452
25.3034 | -106.6484
25.3071 | -106.6527
25.3107 | -106.6542
25.3144 | -106.6564
25.318 | -106.6613
25.3217 | -106.6654
25.3253 | -106.6693
25.329 | -106.6774
25.3327 | -106.6816
25.3363 | -106.6883
25.34 | -106.6968
25.3436 | -106.7055
25.3473 | -106.7132
25.351 | -106.7215
25.3546 | -106.7318
25.3583 | -106.7404
25.3619 | -106.7492
25.3656 | -106.7586
25.3692 | -106.7682
25.3729 | -106.7765
25.3766 | -106.7862
25.3802 | -106.7983
25.3839 | -106.8088
25.3875 | -106.8186
25.3912 | -106.83
25.3949 | -106.8395
25.3985 | -106.8474
25.4022 | -106.8566
25.4058 | -106.8668
25.4095 | -106.8761
25.4131 | -106.8851
25.4168 | -106.8945
25.4205 | -106.9038
25.4241 | -106.9145
25.4278 | -106.9236
25.4314 | -106.9324
25.4351 | -106.9405
25.4387 | -106.9484
25.4424 | -106.9584
25.4461 | -106.9675
25.4497 | -106.9756
25.4534 | -106.9822
25.457 | -106.9884
25.4607 | -106.994
25.4644 | -106.9997
25.468 | -107.0049
25.4717 | -107.008
25.4753 | -107.0122
25.479 | -107.017
25.4826 | -107.0202
25.4863 | -107.0225
25.49 | -107.0227
25.4936 | -107.0218
25.4973 | -107.0209
25.5009 | -107.0198
25.5046 | -107.0199
25.5083 | -107.0201
25.5119 | -107.0202
25.5156 | -107.0174
25.5192 | -107.0151
25.5229 | -107.0125
25.5265 | -107.0084
25.5302 | -107.0056
25.5339 | -107.0012
25.5375 | -106.9945
25.5412 | -106.9882
25.5448 | -106.9814
25.5485 | -106.9754
25.5522 | -106.9687
25.5558 | -106.9613
25.5595 | -106.9528
25.5631 | -106.9444
25.5668 | -106.9374
25.5704 | -106.9289
25.5741 | -106.9198
25.5778 | -106.9117
25.5814 | -106.9035
25.5851 | -106.895
25.5887 | -106.8865
25.5924 | -106.8771
25.596 | -106.8654
25.5997 | -106.8543
25.6034 | -106.8438
25.607 | -106.833
25.6107 | -106.8219
25.6143 | -106.8124
25.618 | -106.8036
25.6217 | -106.7959
25.6253 | -106.7873
25.629 | -106.7793
25.6326 | -106.7709
25.6363 | -106.7618
25.6399 | -106.7516
25.6436 | -106.7425
25.6473 | -106.734
25.6509 | -106.7266
25.6546 | -106.7194
25.6582 | -106.713
25.6619 | -106.7051
25.6656 | -106.6997
25.6692 | -106.6967
25.6729 | -106.6925
25.6765 | -106.6875
25.6802 | -106.6826
25.6838 | -106.6786
25.6875 | -106.6747
25.6912 | -106.6705
25.6948 | -106.6667
25.6985 | -106.6622
25.7021 | -106.6574
25.7058 | -106.6523
25.7095 | -106.6496
25.7131 | -106.6483
25.7168 | -106.6455
25.7204 | -106.6458
25.7241 | -106.6449
25.7277 | -106.6433
25.7314 | -106.6425
25.7351 | -106.6431
25.7387 | -106.6442
25.7424 | -106.6441
25.746 | -106.6449
25.7497 | -106.6464
25.7533 | -106.6468
25.757 | -106.6491
25.7607 | -106.6536
25.7643 | -106.6561
25.768 | -106.6575
25.7716 | -106.6601
25.7753 | -106.6624
25.779 | -106.6654
25.7826 | -106.6691
25.7863 | -106.6732
25.7899 | -106.6771
25.7936 | -106.6794
25.7972 | -106.6833
25.8009 | -106.6862
25.8046 | -106.689
25.8082 | -106.6905
25.8119 | -106.6945
25.8155 | -106.6988
25.8192 | -106.7023
25.8229 | -106.7063
25.8265 | -106.7105
25.8302 | -106.7151
25.8338 | -106.7198
25.8375 | -106.7244
25.8411 | -106.7288
25.8448 | -106.7347
25.8485 | -106.7395
25.8521 | -106.7451
25.8558 | -106.7508
25.8594 | -106.7511
25.8631 | -106.7524
25.8668 | -106.7526
25.8704 | -106.752
25.8741 | -106.7516
25.8777 | -106.751
25.8814 | -106.7519
25.885 | -106.7518
25.8887 | -106.7515
25.8924 | -106.7525
25.896 | -106.7518
25.8997 | -106.7512
25.9033 | -106.7526
25.907 | -106.7529
25.9107 | -106.7528
25.9143 | -106.7508
25.918 | -106.75
25.9216 | -106.7483
25.9253 | -106.7466
25.9289 | -106.7447
25.9326 | -106.7414
25.9363 | -106.7402
25.9399 | -106.739
25.9436 | -106.7371
25.9472 | -106.7348
25.9509 | -106.7318
25.9545 | -106.7296
25.9582 | -106.7255
25.9619 | -106.7218
25.9655 | -106.7198
25.9692 | -106.7175
25.9728 | -106.7147
25.9765 | -106.7117
25.9802 | -106.7072
25.9838 | -106.7031
25.9875 | -106.7005
25.9911 | -106.6987
25.9948 | -106.6952
25.9984 | -106.693
26.0021 | -106.6923
26.0058 | -106.6933
26.0094 | -106.6928
26.0131 | -106.6924
26.0167 | -106.6922
26.0204 | -106.6913
26.0241 | -106.69
26.0277 | -106.689
26.0314 | -106.6854
26.035 | -106.6833
26.0387 | -106.6825
26.0423 | -106.6819
26.046 | -106.6821
26.0497 | -106.6813
26.0533 | -106.6809
26.057 | -106.6812
26.0606 | -106.6809
26.0643 | -106.6807
26.0679 | -106.68
26.0716 | -106.6795
26.0753 | -106.6774
26.0789 | -106.6755
26.0826 | -106.6733
26.0862 | -106.6702
26.0899 | -106.6677
26.0936 | -106.6647
26.0972 | -106.6612
26.1009 | -106.6583
26.1045 | -106.6549
26.1082 | -106.6504
26.1118 | -106.6472
26.1155 | -106.6429
26.1192 | -106.639
26.1228 | -106.6336
26.1265 | -106.6277
26.1301 | -106.6234
26.1338 | -106.6188
26.1375 | -106.614
26.1411 | -106.6091
26.1448 | -106.6057
26.1484 | -106.6001
26.1521 | -106.5924
26.1557 | -106.5865
26.1594 | -106.5831
26.1631 | -106.5784
26.1667 | -106.5725
26.1704 | -106.5672
26.174 | -106.5636
26.1777 | -106.5576
26.1814 | -106.5519
26.185 | -106.5465
26.1887 | -106.5411
26.1923 | -106.5377
26.196 | -106.5341
26.1996 | -106.5293
26.2033 | -106.5237
26.207 | -106.5189
26.2106 | -106.5135
26.2143 | -106.5094
26.2179 | -106.5056
26.2216 | -106.5035
26.2252 | -106.5021
26.2289 | -106.5026
26.2326 | -106.5038
26.2362 | -106.505
26.2399 | -106.5061
26.2435 | -106.5064
26.2472 | -106.5046
26.2509 | -106.5025
26.2545 | -106.5002
26.2582 | -106.4983
26.2618 | -106.4972
26.2655 | -106.4961
26.2691 | -106.4936
26.2728 | -106.4911
26.2765 | -106.4902
26.2801 | -106.4889
26.2838 | -106.4897
26.2874 | -106.4895
26.2911 | -106.4893
26.2948 | -106.4891
26.2984 | -106.4892
26.3021 | -106.488
26.3057 | -106.4847
26.3094 | -106.4818
26.313 | -106.4798
26.3167 | -106.4786
26.3204 | -106.4782
26.324 | -106.4748
26.3277 | -106.4728
26.3313 | -106.4724
26.335 | -106.47
26.3387 | -106.4678
26.3423 | -106.4652
26.346 | -106.4621
26.3496 | -106.4608
26.3533 | -106.4578
26.3569 | -106.4522
26.3606 | -106.4486
26.3643 | -106.4462
26.3679 | -106.4423
26.3716 | -106.4376
26.3752 | -106.4314
26.3789 | -106.4255
26.3826 | -106.4194
26.3862 | -106.412
26.3899 | -106.4061
26.3935 | -106.4012
26.3972 | -106.3953
26.4008 | -106.3904
26.4045 | -106.3863
26.4082 | -106.3811
26.4118 | -106.3757
26.4155 | -106.3703
26.4191 | -106.3644
26.4228 | -106.3591
26.4264 | -106.3543
26.4301 | -106.3498
26.4338 | -106.3444
26.4374 | -106.3398
26.4411 | -106.3347
26.4447 | -106.3303
26.4484 | -106.326
26.4521 | -106.3204
26.4557 | -106.3169
26.4594 | -106.3138
26.463 | -106.3101
26.4667 | -106.3078
26.4703 | -106.3066
26.474 | -106.3041
26.4777 | -106.3013
26.4813 | -106.2999
26.485 | -106.2974
26.4886 | -106.2962
26.4923 | -106.2943
26.496 | -106.2932
26.4996 | -106.2926
26.5033 | -106.2918
26.5069 | -106.2917
26.5106 | -106.2914
26.5142 | -106.2931
26.5179 | -106.293
26.5216 | -106.2953
26.5252 | -106.2994
26.5289 | -106.2991
26.5325 | -106.2993
26.5362 | -106.303
26.5399 | -106.3068
26.5435 | -106.3078
26.5472 | -106.3104
26.5508 | -106.3148
26.5545 | -106.3192
26.5581 | -106.3234
26.5618 | -106.3281
26.5655 | -106.3317
26.5691 | -106.3352
26.5728 | -106.3395
26.5764 | -106.3446
26.5801 | -106.3501
26.5837 | -106.3557
26.5874 | -106.3619
26.5911 | -106.3682
26.5947 | -106.371
26.5984 | -106.3711
26.602 | -106.3736
26.6057 | -106.3772
26.6094 | -106.3815
26.613 | -106.3851
26.6167 | -106.3889
26.6203 | -106.3947
26.624 | -106.3997
26.6276 | -106.4043
26.6313 | -106.4079
26.635 | -106.4098
26.6386 | -106.413
26.6423 | -106.417
26.6459 | -106.4209
26.6496 | -106.4235
26.6533 | -106.4266
26.6569 | -106.4298
26.6606 | -106.4327
26.6642 | -106.4363
26.6679 | -106.4403
26.6715 | -106.4427
26.6752 | -106.4465
26.6789 | -106.4499
26.6825 | -106.4519
26.6862 | -106.4538
26.6898 | -106.4554
26.6935 | -106.4562
26.6972 | -106.4591
26.7008 | -106.4623
26.7045 | -106.464
26.7081 | -106.4639
26.7118 | -106.463
26.7154 | -106.4632
26.7191 | -106.4648
26.7228 | -106.4664
26.7264 | -106.4673
26.7301 | -106.4684
26.7337 | -106.4674
26.7374 | -106.4669
26.741 | -106.4693
26.7447 | -106.4707
26.7484 | -106.4725
26.752 | -106.475
26.7557 | -106.4785
26.7593 | -106.482
26.763 | -106.484
26.7667 | -106.4835
26.7703 | -106.4831
26.774 | -106.4848
26.7776 | -106.4835
26.7813 | -106.4831
26.7849 | -106.4842
26.7886 | -106.485
26.7923 | -106.4855
26.7959 | -106.486
26.7996 | -106.4878
26.8032 | -106.4891
26.8069 | -106.4896
26.8106 | -106.4898
26.8142 | -106.489
26.8179 | -106.4906
26.8215 | -106.4926
26.8252 | -106.4937
26.8288 | -106.4958
26.8325 | -106.4964
26.8362 | -106.496
26.8398 | -106.4973
26.8435 | -106.4986
26.8471 | -106.4977
26.8508 | -106.4974
26.8545 | -106.4983
26.8581 | -106.4986
26.8618 | -106.4991
26.8654 | -106.499
26.8691 | -106.499
26.8727 | -106.4993
26.8764 | -106.4996
26.8801 | -106.4996
26.8837 | -106.4974
26.8874 | -106.4949
26.891 | -106.4932
26.8947 | -106.4928
26.8983 | -106.4919
26.902 | -106.4915
26.9057 | -106.4949
26.9093 | -106.4961
26.913 | -106.4929
26.9166 | -106.4903
26.9203 | -106.4894
26.924 | -106.4898
26.9276 | -106.4881
26.9313 | -106.4868
26.9349 | -106.4866
26.9386 | -106.4858
26.9422 | -106.4844
26.9459 | -106.4833
26.9496 | -106.483
26.9532 | -106.4833
26.9569 | -106.4834
26.9605 | -106.4818
26.9642 | -106.4808
26.9679 | -106.4789
26.9715 | -106.4778
26.9752 | -106.481
26.9788 | -106.484
26.9825 | -106.4836
26.9861 | -106.4842
26.9898 | -106.4844
26.9935 | -106.4842
26.9971 | -106.4837
27.0008 | -106.4843
27.0044 | -106.4851
27.0081 | -106.4857
27.0118 | -106.4873
27.0154 | -106.4897
27.0191 | -106.4896
27.0227 | -106.4893
27.0264 | -106.4892
27.03 | -106.4895
27.0337 | -106.4893
27.0374 | -106.4882
27.041 | -106.4876
27.0447 | -106.4874
27.0483 | -106.4858
27.052 | -106.4864
27.0556 | -106.4861
27.0593 | -106.4849
27.063 | -106.4841
27.0666 | -106.483
27.0703 | -106.4819
27.0739 | -106.4818
27.0776 | -106.4814
27.0813 | -106.4795
27.0849 | -106.478
27.0886 | -106.4777
27.0922 | -106.4769
27.0959 | -106.4764
27.0995 | -106.4747
27.1032 | -106.4711
27.1069 | -106.4677
27.1105 | -106.4654
27.1142 | -106.464
27.1178 | -106.4608
27.1215 | -106.4575
27.1252 | -106.4552
27.1288 | -106.4527
27.1325 | -106.4509
27.1361 | -106.4487
27.1398 | -106.4455
27.1434 | -106.4418
27.1471 | -106.4408
27.1508 | -106.4393
27.1544 | -106.4381
27.1581 | -106.4388
27.1617 | -106.4392
27.1654 | -106.437
27.1691 | -106.4353
27.1727 | -106.4345
27.1764 | -106.4343
27.18 | -106.4329
27.1837 | -106.4304
27.1873 | -106.4292
27.191 | -106.4288
27.1947 | -106.4293
27.1983 | -106.4283
27.202 | -106.4272
27.2056 | -106.4279
27.2093 | -106.4275
27.2129 | -106.4267
27.2166 | -106.4263
27.2203 | -106.4256
27.2239 | -106.4244
27.2276 | -106.4254
27.2312 | -106.4263
27.2349 | -106.4257
27.2386 | -106.4264
27.2422 | -106.4283
27.2459 | -106.4286
27.2495 | -106.4281
27.2532 | -106.4289
27.2568 | -106.4285
27.2605 | -106.4267
27.2642 | -106.4269
27.2678 | -106.4283
27.2715 | -106.4271
27.2751 | -106.4265
27.2788 | -106.4239
27.2825 | -106.4192
27.2861 | -106.4156
27.2898 | -106.4127
27.2934 | -106.4121
27.2971 | -106.4109
27.3007 | -106.4088
27.3044 | -106.4057
27.3081 | -106.4019
27.3117 | -106.3983
27.3154 | -106.3939
27.319 | -106.3887
27.3227 | -106.3848
27.3264 | -106.3805
27.33 | -106.3757
27.3337 | -106.3688
27.3373 | -106.3626
27.341 | -106.3576
27.3446 | -106.3536
27.3483 | -106.3499
27.352 | -106.3457
27.3556 | -106.3393
27.3593 | -106.3309
27.3629 | -106.325
27.3666 | -106.3196
27.3702 | -106.3114
27.3739 | -106.3038
27.3776 | -106.2984
27.3812 | -106.2937
27.3849 | -106.288
27.3885 | -106.2826
27.3922 | -106.278
27.3959 | -106.2734
27.3995 | -106.2682
27.4032 | -106.2651
27.4068 | -106.2615
27.4105 | -106.2562
27.4141 | -106.2502
27.4178 | -106.2452
27.4215 | -106.2401
27.4251 | -106.237
27.4288 | -106.2336
27.4324 | -106.2311
27.4361 | -106.2286
27.4398 | -106.2325
27.4434 | -106.2312
27.4471 | -106.2304
27.4507 | -106.2299
27.4544 | -106.2287
27.458 | -106.2277
27.4617 | -106.2268
27.4654 | -106.2252
27.469 | -106.2256
27.4727 | -106.2265
27.4763 | -106.2278
27.48 | -106.2292
27.4837 | -106.2308
27.4873 | -106.2324
27.491 | -106.2347
27.4946 | -106.2389
27.4983 | -106.2418
27.5019 | -106.2443
27.5056 | -106.2472
27.5093 | -106.2494
27.5129 | -106.2525
27.5166 | -106.2568
27.5202 | -106.2603
27.5239 | -106.2624
27.5276 | -106.2644
27.5312 | -106.2688
27.5349 | -106.2744
27.5385 | -106.2787
27.5422 | -106.2818
27.5458 | -106.2846
27.5495 | -106.2878
27.5532 | -106.2913
27.5568 | -106.2946
27.5605 | -106.2983
27.5641 | -106.3005
27.5678 | -106.3027
27.5714 | -106.3059
27.5751 | -106.3093
27.5788 | -106.3129
27.5824 | -106.3153
27.5861 | -106.3166
27.5897 | -106.3188
27.5934 | -106.3201
27.5971 | -106.3216
27.6007 | -106.3246
27.6044 | -106.3277
27.608 | -106.3288
27.6117 | -106.3289
27.6153 | -106.3291
27.619 | -106.3299
27.6227 | -106.3306
27.6263 | -106.3313
27.63 | -106.3306
27.6336 | -106.3294
27.6373 | -106.329
27.641 | -106.328
27.6446 | -106.3287
27.6483 | -106.3303
27.6519 | -106.3309
27.6556 | -106.3299
27.6592 | -106.3304
27.6629 | -106.3303
27.6666 | -106.3298
27.6702 | -106.3295
27.6739 | -106.3284
27.6775 | -106.3275
27.6812 | -106.3264
27.6849 | -106.3253
27.6885 | -106.3245
27.6922 | -106.3245
27.6958 | -106.3239
27.6995 | -106.3243
27.7031 | -106.3252
27.7068 | -106.3256
27.7105 | -106.3257
27.7141 | -106.3257
27.7178 | -106.3261
27.7214 | -106.3283
27.7251 | -106.3298
27.7287 | -106.3297
27.7324 | -106.3299
27.7361 | -106.3289
27.7397 | -106.3305
27.7434 | -106.3329
27.747 | -106.3349
27.7507 | -106.3341
27.7544 | -106.3351
27.758 | -106.3383
27.7617 | -106.3408
27.7653 | -106.3411
27.769 | -106.3426
27.7726 | -106.3455
27.7763 | -106.3476
27.78 | -106.3482
27.7836 | -106.3479
27.7873 | -106.3497
27.7909 | -106.3513
27.7946 | -106.3538
27.7983 | -106.356
27.8019 | -106.3596
27.8056 | -106.3637
27.8092 | -106.3671
27.8129 | -106.3676
27.8165 | -106.3687
27.8202 | -106.3706
27.8239 | -106.372
27.8275 | -106.3727
27.8312 | -106.3724
27.8348 | -106.3718
27.8385 | -106.3729
27.8422 | -106.375
27.8458 | -106.3772
27.8495 | -106.3776
27.8531 | -106.3757
27.8568 | -106.3704
27.8604 | -106.3622
27.8641 | -106.3558
27.8678 | -106.3521
27.8714 | -106.3492
27.8751 | -106.3446
27.8787 | -106.34
27.8824 | -106.3362
27.886 | -106.3306
27.8897 | -106.3266
27.8934 | -106.3213
27.897 | -106.3166
27.9007 | -106.3116
27.9043 | -106.305
27.908 | -106.2964
27.9117 | -106.2879
27.9153 | -106.2805
27.919 | -106.2722
27.9226 | -106.2624
27.9263 | -106.2376
27.9299 | -106.1992
27.9336 | -106.1701
27.9373 | -106.1485
27.9409 | -106.1289
27.9446 | -106.1086
27.9482 | -106.0886
27.9519 | -106.068
27.9556 | -106.0479
27.9592 | -106.0275
27.9629 | -106.0053
27.9665 | -105.9807
27.9702 | -105.9484
27.9738 | -105.9226
27.9775 | -105.8968
27.9812 | -105.8715
27.9848 | -105.8431
27.9885 | -105.8141
27.9921 | -105.7823
27.9958 | -105.745
27.9995 | -105.7088
28.0031 | -105.6771