## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 266 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 250
Hood | 02-92B/P | 197
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1012.2 | hPa |  | PRC PRO 01
Humidity | 55 | %/Hr
Internal temperature of calibration bench | 23.99 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1017 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.5 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.03 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.99 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3112 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 12/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 18/07/2018
Technical Manager | Anthony HUE | Approval |  | 24/07/2018
CEO | Robert GLEMOT | Approval |  | 24/07/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4777 | 57.544
0.012209 | -36.0447 | 38.4207
0.018313 | -34.9835 | 28.265
0.024417 | -34.5162 | 21.9926
0.030521 | -34.4014 | 18.22
0.036626 | -34.2565 | 15.1599
0.04273 | -34.1603 | 12.7968
0.048834 | -34.1136 | 11.2673
0.054938 | -34.05 | 10.1281
0.061043 | -33.9907 | 9.0513
0.067147 | -33.9802 | 8.2266
0.073251 | -33.9806 | 7.4648
0.079355 | -33.9977 | 6.8029
0.08546 | -34.0109 | 6.2114
0.091564 | -34.0061 | 5.77
0.097668 | -33.9884 | 5.5128
0.10377 | -33.9671 | 5.2604
0.10988 | -33.9505 | 4.9807
0.11598 | -33.9417 | 4.6881
0.12209 | -33.9434 | 4.4051
0.12819 | -33.953 | 4.1235
0.13429 | -33.9636 | 3.8385
0.1404 | -33.9669 | 3.5643
0.1465 | -33.9689 | 3.3099
0.15261 | -33.9663 | 3.092
0.15871 | -33.9601 | 2.93
0.16482 | -33.9522 | 2.8254
0.17092 | -33.9499 | 2.782
0.17702 | -33.9478 | 2.7071
0.18313 | -33.947 | 2.6265
0.18923 | -33.9468 | 2.5168
0.19534 | -33.9457 | 2.3904
0.20144 | -33.9447 | 2.2498
0.20754 | -33.9422 | 2.1308
0.21365 | -33.9414 | 2.0174
0.21975 | -33.9411 | 1.9201
0.22586 | -33.9422 | 1.833
0.23196 | -33.9446 | 1.7602
0.23807 | -33.9484 | 1.6965
0.24417 | -33.9522 | 1.6372
0.25027 | -33.9559 | 1.5744
0.25638 | -33.9581 | 1.5123
0.26248 | -33.9597 | 1.4459
0.26859 | -33.9605 | 1.3811
0.27469 | -33.9608 | 1.311
0.2808 | -33.9617 | 1.2386
0.2869 | -33.9625 | 1.1686
0.293 | -33.9634 | 1.1025
0.29911 | -33.9643 | 1.043
0.30521 | -33.9652 | 0.98448
0.31132 | -33.966 | 0.94519
0.31742 | -33.9672 | 0.902
0.32353 | -33.9676 | 0.86625
0.32963 | -33.9685 | 0.8361
0.33573 | -33.9695 | 0.81428
0.34184 | -33.9709 | 0.79149
0.34794 | -33.9729 | 0.76672
0.35405 | -33.975 | 0.74165
0.36015 | -33.9774 | 0.71547
0.36626 | -33.9794 | 0.68743
0.37236 | -33.9814 | 0.65071
0.37846 | -33.9835 | 0.60699
0.38457 | -33.9843 | 0.55126
0.39067 | -33.9853 | 0.50533
0.39678 | -33.986 | 0.45448
0.40288 | -33.9864 | 0.40159
0.40899 | -33.9864 | 0.3508
0.41509 | -33.9857 | 0.30098
0.42119 | -33.9857 | 0.25835
0.4273 | -33.9857 | 0.21295
0.4334 | -33.9857 | 0.16761
0.43951 | -33.986 | 0.12488
0.44561 | -33.9869 | 0.085582
0.45172 | -33.9884 | 0.053844
0.45782 | -33.9892 | 0.017006
0.46392 | -33.9904 | -0.018948
0.47003 | -33.9915 | -0.05136
0.47613 | -33.9923 | -0.085346
0.48224 | -33.9929 | -0.12
0.48834 | -33.992 | -0.1545
0.49445 | -33.9918 | -0.18959
0.50055 | -33.9913 | -0.22158
0.50665 | -33.9907 | -0.25553
0.51276 | -33.9897 | -0.28927
0.51886 | -33.9888 | -0.32291
0.52497 | -33.9876 | -0.34717
0.53107 | -33.9869 | -0.37648
0.53717 | -33.9863 | -0.4045
0.54328 | -33.9855 | -0.42742
0.54938 | -33.9853 | -0.44997
0.55549 | -33.9854 | -0.47107
0.56159 | -33.9864 | -0.48321
0.5677 | -33.9867 | -0.49819
0.5738 | -33.9878 | -0.51481
0.5799 | -33.9885 | -0.53012
0.58601 | -33.9896 | -0.54083
0.59211 | -33.9912 | -0.54795
0.59822 | -33.9925 | -0.5581
0.60432 | -33.994 | -0.57068
0.61043 | -33.9952 | -0.57934
0.61653 | -33.9964 | -0.59025
0.62263 | -33.998 | -0.60024
0.62874 | -33.9986 | -0.61139
0.63484 | -33.9995 | -0.62129
0.64095 | -34 | -0.63202
0.64705 | -34.0004 | -0.64236
0.65316 | -34.0007 | -0.65527
0.65926 | -34.0008 | -0.67136
0.66536 | -33.9998 | -0.68892
0.67147 | -33.9995 | -0.70244
0.67757 | -33.9992 | -0.71704
0.68368 | -33.9985 | -0.736
0.68978 | -33.9979 | -0.7533
0.69589 | -33.997 | -0.77235
0.70199 | -33.9962 | -0.79494
0.70809 | -33.9954 | -0.81416
0.7142 | -33.9947 | -0.8346
0.7203 | -33.9938 | -0.85377
0.72641 | -33.9932 | -0.87616
0.73251 | -33.9926 | -0.89621
0.73862 | -33.9928 | -0.91792
0.74472 | -33.9926 | -0.93798
0.75082 | -33.9924 | -0.96002
0.75693 | -33.9927 | -0.98179
0.76303 | -33.9933 | -1.0037
0.76914 | -33.995 | -1.0242
0.77524 | -33.996 | -1.0438
0.78135 | -33.9972 | -1.0639
0.78745 | -33.9985 | -1.0839
0.79355 | -33.9997 | -1.1037
0.79966 | -34.0014 | -1.1231
0.80576 | -34.0036 | -1.1426
0.81187 | -34.0051 | -1.164
0.81797 | -34.007 | -1.1851
0.82408 | -34.0086 | -1.2055
0.83018 | -34.0103 | -1.2259
0.83628 | -34.0121 | -1.2443
0.84239 | -34.0132 | -1.2646
0.84849 | -34.0146 | -1.286
0.8546 | -34.016 | -1.3055
0.8607 | -34.0172 | -1.326
0.86681 | -34.0183 | -1.3448
0.87291 | -34.0192 | -1.3634
0.87901 | -34.0197 | -1.3811
0.88512 | -34.0204 | -1.3991
0.89122 | -34.0211 | -1.4171
0.89733 | -34.0215 | -1.4327
0.90343 | -34.0219 | -1.4501
0.90953 | -34.0221 | -1.4644
0.91564 | -34.0221 | -1.4802
0.92174 | -34.0223 | -1.4951
0.92785 | -34.0224 | -1.5096
0.93395 | -34.0225 | -1.5235
0.94006 | -34.0227 | -1.5364
0.94616 | -34.0225 | -1.5483
0.95226 | -34.0226 | -1.5611
0.95837 | -34.0227 | -1.5749
0.96447 | -34.0229 | -1.5878
0.97058 | -34.0228 | -1.6012
0.97668 | -34.0226 | -1.6146
0.98279 | -34.0223 | -1.6274
0.98889 | -34.0222 | -1.6382
0.99499 | -34.0221 | -1.6516
1.0011 | -34.0221 | -1.6624
1.0072 | -34.0217 | -1.6742
1.0133 | -34.0214 | -1.6851
1.0194 | -34.0211 | -1.6982
1.0255 | -34.0208 | -1.7119
1.0316 | -34.0203 | -1.7239
1.0377 | -34.0199 | -1.7376
1.0438 | -34.0194 | -1.7504
1.0499 | -34.0187 | -1.7656
1.056 | -34.018 | -1.7799
1.0621 | -34.0175 | -1.7942
1.0682 | -34.017 | -1.8068
1.0743 | -34.0164 | -1.8222
1.0805 | -34.016 | -1.8366
1.0866 | -34.0151 | -1.8517
1.0927 | -34.0144 | -1.8662
1.0988 | -34.0138 | -1.8811
1.1049 | -34.013 | -1.8953
1.111 | -34.0123 | -1.9107
1.1171 | -34.0117 | -1.9253
1.1232 | -34.0113 | -1.9405
1.1293 | -34.0109 | -1.9546
1.1354 | -34.0106 | -1.9696
1.1415 | -34.0103 | -1.985
1.1476 | -34.0097 | -2.0005
1.1537 | -34.0093 | -2.0152
1.1598 | -34.0094 | -2.0299
1.1659 | -34.0092 | -2.0433
1.172 | -34.009 | -2.0573
1.1781 | -34.0088 | -2.0712
1.1842 | -34.0087 | -2.0854
1.1903 | -34.009 | -2.0964
1.1964 | -34.0089 | -2.1098
1.2025 | -34.0091 | -2.1237
1.2086 | -34.0094 | -2.1363
1.2147 | -34.0094 | -2.1488
1.2209 | -34.0098 | -2.1601
1.227 | -34.0105 | -2.1709
1.2331 | -34.0109 | -2.183
1.2392 | -34.0112 | -2.1948
1.2453 | -34.0117 | -2.2059
1.2514 | -34.0121 | -2.2159
1.2575 | -34.0126 | -2.2269
1.2636 | -34.0134 | -2.2358
1.2697 | -34.014 | -2.2451
1.2758 | -34.0147 | -2.2562
1.2819 | -34.0154 | -2.2662
1.288 | -34.016 | -2.2769
1.2941 | -34.0168 | -2.2855
1.3002 | -34.0175 | -2.2936
1.3063 | -34.018 | -2.3027
1.3124 | -34.0186 | -2.3117
1.3185 | -34.0193 | -2.3208
1.3246 | -34.02 | -2.3299
1.3307 | -34.0208 | -2.3385
1.3368 | -34.0216 | -2.3472
1.3429 | -34.0222 | -2.3562
1.349 | -34.0227 | -2.3655
1.3551 | -34.0234 | -2.3748
1.3613 | -34.0239 | -2.3831
1.3674 | -34.0245 | -2.3931
1.3735 | -34.025 | -2.4029
1.3796 | -34.0254 | -2.4119
1.3857 | -34.0257 | -2.4212
1.3918 | -34.0263 | -2.4314
1.3979 | -34.0268 | -2.4406
1.404 | -34.0273 | -2.4519
1.4101 | -34.0277 | -2.4622
1.4162 | -34.0281 | -2.4728
1.4223 | -34.0286 | -2.4841
1.4284 | -34.0291 | -2.4942
1.4345 | -34.0293 | -2.5057
1.4406 | -34.0296 | -2.5191
1.4467 | -34.0301 | -2.5306
1.4528 | -34.0305 | -2.542
1.4589 | -34.031 | -2.5538
1.465 | -34.0312 | -2.5647
1.4711 | -34.0317 | -2.576
1.4772 | -34.0322 | -2.5893
1.4833 | -34.0327 | -2.6022
1.4894 | -34.0332 | -2.6144
1.4955 | -34.0338 | -2.6263
1.5016 | -34.0343 | -2.6387
1.5078 | -34.0347 | -2.6525
1.5139 | -34.0351 | -2.6655
1.52 | -34.0355 | -2.6781
1.5261 | -34.0358 | -2.6907
1.5322 | -34.0362 | -2.7046
1.5383 | -34.0366 | -2.7171
1.5444 | -34.0371 | -2.7298
1.5505 | -34.0376 | -2.7431
1.5566 | -34.0381 | -2.7561
1.5627 | -34.0385 | -2.7698
1.5688 | -34.0391 | -2.7833
1.5749 | -34.0395 | -2.796
1.581 | -34.0398 | -2.8094
1.5871 | -34.0402 | -2.823
1.5932 | -34.0405 | -2.8359
1.5993 | -34.0409 | -2.85
1.6054 | -34.0413 | -2.8632
1.6115 | -34.0417 | -2.8773
1.6176 | -34.042 | -2.8899
1.6237 | -34.0422 | -2.9035
1.6298 | -34.0425 | -2.9161
1.6359 | -34.0429 | -2.9289
1.642 | -34.0433 | -2.941
1.6482 | -34.0435 | -2.954
1.6543 | -34.0438 | -2.9677
1.6604 | -34.0442 | -2.9798
1.6665 | -34.0443 | -2.9923
1.6726 | -34.0445 | -3.0049
1.6787 | -34.0448 | -3.0176
1.6848 | -34.0448 | -3.0303
1.6909 | -34.045 | -3.0431
1.697 | -34.0453 | -3.0554
1.7031 | -34.0453 | -3.0681
1.7092 | -34.0453 | -3.0804
1.7153 | -34.0455 | -3.0925
1.7214 | -34.0453 | -3.1048
1.7275 | -34.0453 | -3.1175
1.7336 | -34.0453 | -3.1304
1.7397 | -34.0454 | -3.1429
1.7458 | -34.0454 | -3.1552
1.7519 | -34.0455 | -3.1674
1.758 | -34.0455 | -3.1795
1.7641 | -34.0457 | -3.1912
1.7702 | -34.0457 | -3.2042
1.7763 | -34.0456 | -3.2158
1.7824 | -34.0456 | -3.2284
1.7885 | -34.0456 | -3.2408
1.7947 | -34.0458 | -3.2529
1.8008 | -34.0458 | -3.2643
1.8069 | -34.0458 | -3.276
1.813 | -34.0459 | -3.2883
1.8191 | -34.0459 | -3.301
1.8252 | -34.046 | -3.3137
1.8313 | -34.0462 | -3.326
1.8374 | -34.0463 | -3.3378
1.8435 | -34.0464 | -3.3498
1.8496 | -34.0465 | -3.3628
1.8557 | -34.0467 | -3.3755
1.8618 | -34.047 | -3.3876
1.8679 | -34.0471 | -3.3996
1.874 | -34.0472 | -3.4116
1.8801 | -34.0474 | -3.4239
1.8862 | -34.0477 | -3.4359
1.8923 | -34.0478 | -3.448
1.8984 | -34.0479 | -3.4599
1.9045 | -34.048 | -3.4715
1.9106 | -34.0481 | -3.4836
1.9167 | -34.0482 | -3.4957
1.9228 | -34.0483 | -3.5075
1.9289 | -34.0485 | -3.5192
1.9351 | -34.0487 | -3.53
1.9412 | -34.0489 | -3.5412
1.9473 | -34.0492 | -3.5531
1.9534 | -34.0494 | -3.5645
1.9595 | -34.0496 | -3.5762
1.9656 | -34.0499 | -3.5883
1.9717 | -34.0501 | -3.5999
1.9778 | -34.0505 | -3.6112
1.9839 | -34.0507 | -3.6226
1.99 | -34.0509 | -3.6335
1.9961 | -34.051 | -3.6456
2.0022 | -34.0513 | -3.6568
2.0083 | -34.0514 | -3.668
2.0144 | -34.0516 | -3.6793
2.0205 | -34.0519 | -3.6909
2.0266 | -34.0523 | -3.7022
2.0327 | -34.0525 | -3.7136
2.0388 | -34.0529 | -3.7251
2.0449 | -34.0531 | -3.7363
2.051 | -34.0534 | -3.7476
2.0571 | -34.0537 | -3.7589
2.0632 | -34.0539 | -3.7703
2.0693 | -34.0541 | -3.7823
2.0754 | -34.0544 | -3.7935
2.0816 | -34.0547 | -3.8055
2.0877 | -34.0549 | -3.817
2.0938 | -34.0552 | -3.8289
2.0999 | -34.0554 | -3.8406
2.106 | -34.0556 | -3.8521
2.1121 | -34.0558 | -3.864
2.1182 | -34.056 | -3.8747
2.1243 | -34.0563 | -3.8854
2.1304 | -34.0565 | -3.8965
2.1365 | -34.0568 | -3.9081
2.1426 | -34.057 | -3.92
2.1487 | -34.0573 | -3.9312
2.1548 | -34.0577 | -3.9429
2.1609 | -34.0581 | -3.9543
2.167 | -34.0585 | -3.9654
2.1731 | -34.0588 | -3.9773
2.1792 | -34.059 | -3.9883
2.1853 | -34.0593 | -4.0001
2.1914 | -34.0596 | -4.0114
2.1975 | -34.06 | -4.0228
2.2036 | -34.0602 | -4.0348
2.2097 | -34.0604 | -4.0461
2.2158 | -34.0608 | -4.0567
2.222 | -34.061 | -4.0672
2.2281 | -34.0612 | -4.0785
2.2342 | -34.0615 | -4.0897
2.2403 | -34.0617 | -4.1004
2.2464 | -34.062 | -4.1111
2.2525 | -34.0622 | -4.1225
2.2586 | -34.0625 | -4.133
2.2647 | -34.0628 | -4.1435
2.2708 | -34.0631 | -4.1547
2.2769 | -34.0634 | -4.1653
2.283 | -34.0636 | -4.1764
2.2891 | -34.0638 | -4.1873
2.2952 | -34.064 | -4.1978
2.3013 | -34.0642 | -4.2086
2.3074 | -34.0644 | -4.2191
2.3135 | -34.0646 | -4.2304
2.3196 | -34.0648 | -4.2406
2.3257 | -34.065 | -4.2512
2.3318 | -34.0652 | -4.2621
2.3379 | -34.0653 | -4.2729
2.344 | -34.0655 | -4.2842
2.3501 | -34.0657 | -4.2949
2.3562 | -34.0658 | -4.3054
2.3623 | -34.0659 | -4.3157
2.3685 | -34.066 | -4.3263
2.3746 | -34.0662 | -4.3375
2.3807 | -34.0664 | -4.3479
2.3868 | -34.0666 | -4.3586
2.3929 | -34.0667 | -4.3693
2.399 | -34.0668 | -4.3797
2.4051 | -34.0671 | -4.3902
2.4112 | -34.0672 | -4.4008
2.4173 | -34.0674 | -4.4114
2.4234 | -34.0675 | -4.4222
2.4295 | -34.0676 | -4.4331
2.4356 | -34.0678 | -4.4439
2.4417 | -34.0679 | -4.455
2.4478 | -34.068 | -4.4658
2.4539 | -34.0682 | -4.4761
2.46 | -34.0683 | -4.4875
2.4661 | -34.0685 | -4.4986
2.4722 | -34.0686 | -4.5088
2.4783 | -34.0687 | -4.5192
2.4844 | -34.0688 | -4.5303
2.4905 | -34.0689 | -4.5412
2.4966 | -34.0691 | -4.5519
2.5027 | -34.0692 | -4.5624
2.5089 | -34.0693 | -4.5733
2.515 | -34.0693 | -4.5841
2.5211 | -34.0694 | -4.5949
2.5272 | -34.0695 | -4.6058
2.5333 | -34.0697 | -4.6168
2.5394 | -34.0699 | -4.6275
2.5455 | -34.07 | -4.6389
2.5516 | -34.0701 | -4.6493
2.5577 | -34.0703 | -4.6607
2.5638 | -34.0705 | -4.6724
2.5699 | -34.0705 | -4.6838
2.576 | -34.0706 | -4.6954
2.5821 | -34.0707 | -4.7066
2.5882 | -34.0709 | -4.7177
2.5943 | -34.071 | -4.7286
2.6004 | -34.0712 | -4.7396
2.6065 | -34.0713 | -4.75
2.6126 | -34.0715 | -4.7605
2.6187 | -34.0716 | -4.7711
2.6248 | -34.0717 | -4.7819
2.6309 | -34.0719 | -4.7933
2.637 | -34.0722 | -4.8049
2.6431 | -34.0723 | -4.8155
2.6492 | -34.0724 | -4.8268
2.6554 | -34.0727 | -4.8377
2.6615 | -34.0728 | -4.8487
2.6676 | -34.073 | -4.8601
2.6737 | -34.0733 | -4.8713
2.6798 | -34.0736 | -4.8825
2.6859 | -34.0738 | -4.8938
2.692 | -34.074 | -4.9047
2.6981 | -34.0742 | -4.9158
2.7042 | -34.0744 | -4.927
2.7103 | -34.0746 | -4.9376
2.7164 | -34.0749 | -4.9486
2.7225 | -34.0752 | -4.96
2.7286 | -34.0754 | -4.971
2.7347 | -34.0756 | -4.9816
2.7408 | -34.0759 | -4.9918
2.7469 | -34.0761 | -5.003
2.753 | -34.0764 | -5.0137
2.7591 | -34.0767 | -5.025
2.7652 | -34.0769 | -5.0365
2.7713 | -34.0771 | -5.0473
2.7774 | -34.0774 | -5.0582
2.7835 | -34.0776 | -5.0693
2.7896 | -34.0778 | -5.0798
2.7958 | -34.0781 | -5.0906
2.8019 | -34.0784 | -5.1011
2.808 | -34.0786 | -5.1119
2.8141 | -34.079 | -5.1231
2.8202 | -34.0793 | -5.134
2.8263 | -34.0796 | -5.1449
2.8324 | -34.0799 | -5.1563
2.8385 | -34.0803 | -5.1671
2.8446 | -34.0807 | -5.1779
2.8507 | -34.081 | -5.1887
2.8568 | -34.0814 | -5.1992
2.8629 | -34.0816 | -5.2098
2.869 | -34.082 | -5.2209
2.8751 | -34.0823 | -5.2317
2.8812 | -34.0827 | -5.2429
2.8873 | -34.083 | -5.2539
2.8934 | -34.0833 | -5.2649
2.8995 | -34.0836 | -5.276
2.9056 | -34.0839 | -5.2871
2.9117 | -34.0842 | -5.2981
2.9178 | -34.0846 | -5.3088
2.9239 | -34.085 | -5.3203
2.93 | -34.0854 | -5.3317
2.9361 | -34.0857 | -5.3425
2.9423 | -34.086 | -5.3532
2.9484 | -34.0864 | -5.3642
2.9545 | -34.0868 | -5.3752
2.9606 | -34.0872 | -5.3865
2.9667 | -34.0875 | -5.3976
2.9728 | -34.0878 | -5.4088
2.9789 | -34.0883 | -5.4199
2.985 | -34.0887 | -5.431
2.9911 | -34.089 | -5.4423
2.9972 | -34.0895 | -5.4533
3.0033 | -34.0899 | -5.4647
3.0094 | -34.0902 | -5.4758
3.0155 | -34.0906 | -5.4866
3.0216 | -34.091 | -5.4978
3.0277 | -34.0913 | -5.5094
3.0338 | -34.0917 | -5.5207
3.0399 | -34.0921 | -5.532
3.046 | -34.0924 | -5.5434
3.0521 | -34.0927 | -5.5544
3.0582 | -34.093 | -5.5663
3.0643 | -34.0933 | -5.5776
3.0704 | -34.0937 | -5.5893
3.0765 | -34.0939 | -5.6004
3.0827 | -34.0942 | -5.6115
3.0888 | -34.0945 | -5.6227
3.0949 | -34.0949 | -5.6339
3.101 | -34.0953 | -5.6452
3.1071 | -34.0956 | -5.6569
3.1132 | -34.0959 | -5.6685
3.1193 | -34.0963 | -5.6799
3.1254 | -34.0966 | -5.6911
3.1315 | -34.097 | -5.7028
3.1376 | -34.0973 | -5.7141
3.1437 | -34.0976 | -5.7256
3.1498 | -34.0979 | -5.7375
3.1559 | -34.0982 | -5.7489
3.162 | -34.0985 | -5.7608
3.1681 | -34.0987 | -5.7726
3.1742 | -34.099 | -5.784
3.1803 | -34.0993 | -5.7955
3.1864 | -34.0997 | -5.8067
3.1925 | -34.1 | -5.8181
3.1986 | -34.1002 | -5.8299
3.2047 | -34.1005 | -5.8418
3.2108 | -34.1007 | -5.8528
3.2169 | -34.101 | -5.8643
3.223 | -34.1013 | -5.8759
3.2292 | -34.1015 | -5.8876
3.2353 | -34.1017 | -5.8983
3.2414 | -34.1019 | -5.9096
3.2475 | -34.1021 | -5.921
3.2536 | -34.1024 | -5.9323
3.2597 | -34.1027 | -5.9431
3.2658 | -34.1029 | -5.954
3.2719 | -34.1031 | -5.965
3.278 | -34.1033 | -5.9765
3.2841 | -34.1036 | -5.9874
3.2902 | -34.1038 | -5.9988
3.2963 | -34.1041 | -6.0101
3.3024 | -34.1042 | -6.0212
3.3085 | -34.1044 | -6.0326
3.3146 | -34.1046 | -6.0437
3.3207 | -34.1048 | -6.0551
3.3268 | -34.1051 | -6.0665
3.3329 | -34.1052 | -6.0771
3.339 | -34.1054 | -6.0878
3.3451 | -34.1056 | -6.0989
3.3512 | -34.1058 | -6.1096
3.3573 | -34.106 | -6.1207
3.3634 | -34.1062 | -6.1318
3.3696 | -34.1064 | -6.1431
3.3757 | -34.1065 | -6.1538
3.3818 | -34.1068 | -6.1646
3.3879 | -34.1069 | -6.1756
3.394 | -34.1071 | -6.187
3.4001 | -34.1073 | -6.1978
3.4062 | -34.1074 | -6.2086
3.4123 | -34.1075 | -6.2192
3.4184 | -34.1077 | -6.2304
3.4245 | -34.108 | -6.2412
3.4306 | -34.108 | -6.2518
3.4367 | -34.1082 | -6.2624
3.4428 | -34.1084 | -6.2728
3.4489 | -34.1085 | -6.2837
3.455 | -34.1087 | -6.2945
3.4611 | -34.1089 | -6.3057
3.4672 | -34.109 | -6.3167
3.4733 | -34.1092 | -6.3278
3.4794 | -34.1094 | -6.3386
3.4855 | -34.1096 | -6.349
3.4916 | -34.1097 | -6.3594
3.4977 | -34.1099 | -6.3707
3.5038 | -34.1101 | -6.3815
3.5099 | -34.1102 | -6.3918
3.5161 | -34.1104 | -6.4022
3.5222 | -34.1105 | -6.4126
3.5283 | -34.1106 | -6.4234
3.5344 | -34.1107 | -6.4342
3.5405 | -34.1109 | -6.4448
3.5466 | -34.1111 | -6.4557
3.5527 | -34.1112 | -6.4666
3.5588 | -34.1113 | -6.4772
3.5649 | -34.1114 | -6.4879
3.571 | -34.1116 | -6.4986
3.5771 | -34.1117 | -6.5096
3.5832 | -34.1119 | -6.5196
3.5893 | -34.1121 | -6.5302
3.5954 | -34.1123 | -6.5411
3.6015 | -34.1124 | -6.5519
3.6076 | -34.1127 | -6.5625
3.6137 | -34.1128 | -6.5726
3.6198 | -34.1131 | -6.583
3.6259 | -34.1132 | -6.5939
3.632 | -34.1134 | -6.6046
3.6381 | -34.1136 | -6.6153
3.6442 | -34.1137 | -6.626
3.6503 | -34.114 | -6.6362
3.6565 | -34.1142 | -6.647
3.6626 | -34.1143 | -6.6576
3.6687 | -34.1146 | -6.6683
3.6748 | -34.1148 | -6.6787
3.6809 | -34.1151 | -6.6894
3.687 | -34.1153 | -6.7004
3.6931 | -34.1155 | -6.7107
3.6992 | -34.1157 | -6.7214
3.7053 | -34.1159 | -6.732
3.7114 | -34.1161 | -6.7427
3.7175 | -34.1163 | -6.7534
3.7236 | -34.1166 | -6.7635
3.7297 | -34.1168 | -6.7743
3.7358 | -34.117 | -6.7852
3.7419 | -34.1172 | -6.7961
3.748 | -34.1175 | -6.8067
3.7541 | -34.1177 | -6.8173
3.7602 | -34.118 | -6.8279
3.7663 | -34.1182 | -6.8384
3.7724 | -34.1184 | -6.8492
3.7785 | -34.1187 | -6.8596
3.7846 | -34.1189 | -6.8701
3.7907 | -34.1192 | -6.8811
3.7969 | -34.1194 | -6.8918
3.803 | -34.1197 | -6.9021
3.8091 | -34.1199 | -6.9127
3.8152 | -34.1202 | -6.9233
3.8213 | -34.1204 | -6.9341
3.8274 | -34.1207 | -6.9445
3.8335 | -34.121 | -6.9553
3.8396 | -34.1213 | -6.966
3.8457 | -34.1216 | -6.9767
3.8518 | -34.122 | -6.9873
3.8579 | -34.1223 | -6.9979
3.864 | -34.1225 | -7.0087
3.8701 | -34.1229 | -7.0197
3.8762 | -34.1231 | -7.0302
3.8823 | -34.1233 | -7.041
3.8884 | -34.1237 | -7.0517
3.8945 | -34.124 | -7.0626
3.9006 | -34.1243 | -7.0736
3.9067 | -34.1246 | -7.0848
3.9128 | -34.1248 | -7.0955
3.9189 | -34.1251 | -7.106
3.925 | -34.1253 | -7.1169
3.9311 | -34.1256 | -7.1271
3.9372 | -34.1259 | -7.138
3.9434 | -34.1262 | -7.1487
3.9495 | -34.1266 | -7.1595
3.9556 | -34.1269 | -7.1701
3.9617 | -34.1272 | -7.1804
3.9678 | -34.1275 | -7.1908
3.9739 | -34.1278 | -7.2014
3.98 | -34.128 | -7.2122
3.9861 | -34.1283 | -7.2229
3.9922 | -34.1285 | -7.2336
3.9983 | -34.1288 | -7.2439
4.0044 | -34.1291 | -7.2542
4.0105 | -34.1294 | -7.2649
4.0166 | -34.1297 | -7.2756
4.0227 | -34.13 | -7.2862
4.0288 | -34.1302 | -7.2973
4.0349 | -34.1304 | -7.3078
4.041 | -34.1307 | -7.3184
4.0471 | -34.1309 | -7.3291
4.0532 | -34.1312 | -7.3395
4.0593 | -34.1315 | -7.3503
4.0654 | -34.1318 | -7.3607
4.0715 | -34.1321 | -7.3718
4.0776 | -34.1324 | -7.3824
4.0838 | -34.1327 | -7.3935
4.0899 | -34.1329 | -7.4038
4.096 | -34.1332 | -7.4145
4.1021 | -34.1334 | -7.4249
4.1082 | -34.1337 | -7.4353
4.1143 | -34.1339 | -7.4456
4.1204 | -34.1342 | -7.4559
4.1265 | -34.1345 | -7.4663
4.1326 | -34.1347 | -7.4766
4.1387 | -34.1351 | -7.4871
4.1448 | -34.1353 | -7.4978
4.1509 | -34.1355 | -7.5084
4.157 | -34.1358 | -7.5193
4.1631 | -34.136 | -7.5301
4.1692 | -34.1363 | -7.5406
4.1753 | -34.1366 | -7.5511
4.1814 | -34.1367 | -7.5614
4.1875 | -34.137 | -7.5722
4.1936 | -34.1373 | -7.5828
4.1997 | -34.1375 | -7.593
4.2058 | -34.1377 | -7.6038
4.2119 | -34.138 | -7.6142
4.218 | -34.1382 | -7.6245
4.2241 | -34.1384 | -7.6348
4.2303 | -34.1387 | -7.6452
4.2364 | -34.1389 | -7.6557
4.2425 | -34.1391 | -7.6662
4.2486 | -34.1393 | -7.6769
4.2547 | -34.1395 | -7.6871
4.2608 | -34.1397 | -7.6978
4.2669 | -34.1399 | -7.708
4.273 | -34.1401 | -7.7185
4.2791 | -34.1403 | -7.729
4.2852 | -34.1406 | -7.7397
4.2913 | -34.1407 | -7.7499
4.2974 | -34.1409 | -7.7606
4.3035 | -34.1412 | -7.7714
4.3096 | -34.1414 | -7.782
4.3157 | -34.1416 | -7.7923
4.3218 | -34.1418 | -7.8026
4.3279 | -34.142 | -7.8131
4.334 | -34.1422 | -7.8234
4.3401 | -34.1424 | -7.8335
4.3462 | -34.1427 | -7.8441
4.3523 | -34.1429 | -7.8541
4.3584 | -34.1431 | -7.8642
4.3645 | -34.1433 | -7.8745
4.3707 | -34.1435 | -7.8847
4.3768 | -34.1437 | -7.895
4.3829 | -34.1439 | -7.9057
4.389 | -34.1441 | -7.916
4.3951 | -34.1443 | -7.9262
4.4012 | -34.1445 | -7.9369
4.4073 | -34.1447 | -7.9471
4.4134 | -34.1449 | -7.9574
4.4195 | -34.1452 | -7.9677
4.4256 | -34.1454 | -7.9778
4.4317 | -34.1456 | -7.9881
4.4378 | -34.1458 | -7.9981
4.4439 | -34.146 | -8.0081
4.45 | -34.1462 | -8.0184
4.4561 | -34.1463 | -8.0287
4.4622 | -34.1465 | -8.039
4.4683 | -34.1467 | -8.0493
4.4744 | -34.1469 | -8.06
4.4805 | -34.1472 | -8.0699
4.4866 | -34.1474 | -8.0802
4.4927 | -34.1476 | -8.0906
4.4988 | -34.1478 | -8.1009
4.5049 | -34.148 | -8.1108
4.511 | -34.1483 | -8.1208
4.5172 | -34.1485 | -8.131
4.5233 | -34.1487 | -8.1417
4.5294 | -34.1489 | -8.1518
4.5355 | -34.1491 | -8.162
4.5416 | -34.1493 | -8.1723
4.5477 | -34.1496 | -8.1825
4.5538 | -34.1497 | -8.1924
4.5599 | -34.15 | -8.2026
4.566 | -34.1502 | -8.2129
4.5721 | -34.1504 | -8.2231
4.5782 | -34.1507 | -8.2336
4.5843 | -34.1509 | -8.2439
4.5904 | -34.1512 | -8.2542
4.5965 | -34.1514 | -8.2643
4.6026 | -34.1516 | -8.275
4.6087 | -34.1518 | -8.2851
4.6148 | -34.152 | -8.2957
4.6209 | -34.1523 | -8.3058
4.627 | -34.1525 | -8.3163
4.6331 | -34.1527 | -8.3264
4.6392 | -34.1529 | -8.3367
4.6453 | -34.1532 | -8.347
4.6514 | -34.1534 | -8.3573
4.6576 | -34.1536 | -8.3679
4.6637 | -34.1538 | -8.3778
4.6698 | -34.154 | -8.3879
4.6759 | -34.1543 | -8.3986
4.682 | -34.1544 | -8.4088
4.6881 | -34.1546 | -8.419
4.6942 | -34.1549 | -8.4292
4.7003 | -34.1551 | -8.4395
4.7064 | -34.1553 | -8.4496
4.7125 | -34.1556 | -8.4602
4.7186 | -34.1558 | -8.4703
4.7247 | -34.156 | -8.4806
4.7308 | -34.1563 | -8.4909
4.7369 | -34.1565 | -8.5012
4.743 | -34.1567 | -8.5119
4.7491 | -34.157 | -8.5225
4.7552 | -34.1572 | -8.5329
4.7613 | -34.1574 | -8.5433
4.7674 | -34.1576 | -8.5535
4.7735 | -34.1578 | -8.5642
4.7796 | -34.1581 | -8.5748
4.7857 | -34.1584 | -8.5852
4.7918 | -34.1586 | -8.5953
4.7979 | -34.1589 | -8.6056
4.8041 | -34.1591 | -8.6159
4.8102 | -34.1594 | -8.626
4.8163 | -34.1596 | -8.6363
4.8224 | -34.1598 | -8.6466
4.8285 | -34.1601 | -8.6566
4.8346 | -34.1604 | -8.6667
4.8407 | -34.1606 | -8.6772
4.8468 | -34.1608 | -8.6878
4.8529 | -34.1611 | -8.6984
4.859 | -34.1613 | -8.709
4.8651 | -34.1616 | -8.7192
4.8712 | -34.1618 | -8.7294
4.8773 | -34.162 | -8.7398
4.8834 | -34.1623 | -8.7501
4.8895 | -34.1625 | -8.7603
4.8956 | -34.1627 | -8.7706
4.9017 | -34.1629 | -8.7808
4.9078 | -34.1632 | -8.7915
4.9139 | -34.1634 | -8.8015
4.92 | -34.1638 | -8.8116
4.9261 | -34.164 | -8.8216
4.9322 | -34.1643 | -8.8319
4.9383 | -34.1646 | -8.8422
4.9445 | -34.1648 | -8.8526
4.9506 | -34.165 | -8.8627
4.9567 | -34.1653 | -8.8729
4.9628 | -34.1655 | -8.8832
4.9689 | -34.1658 | -8.8939
4.975 | -34.166 | -8.9046
4.9811 | -34.1663 | -8.9146
4.9872 | -34.1666 | -8.9248
4.9933 | -34.1669 | -8.9354
4.9994 | -34.1672 | -8.9461
5.0055 | -34.1675 | -8.9565
5.0116 | -34.1678 | -8.9673
5.0177 | -34.1681 | -8.9779
5.0238 | -34.1684 | -8.9885
5.0299 | -34.1687 | -8.9992
5.036 | -34.169 | -9.0096
5.0421 | -34.1692 | -9.0198
5.0482 | -34.1695 | -9.0302
5.0543 | -34.1698 | -9.0404
5.0604 | -34.1701 | -9.0507
5.0665 | -34.1704 | -9.0609
5.0726 | -34.1707 | -9.0716
5.0787 | -34.171 | -9.0821
5.0848 | -34.1713 | -9.0927
5.091 | -34.1715 | -9.1029
5.0971 | -34.1718 | -9.1129
5.1032 | -34.1721 | -9.1236
5.1093 | -34.1724 | -9.1341
5.1154 | -34.1727 | -9.1442
5.1215 | -34.1729 | -9.1545
5.1276 | -34.1732 | -9.1647
5.1337 | -34.1735 | -9.1748
5.1398 | -34.1738 | -9.185
5.1459 | -34.1741 | -9.1956
5.152 | -34.1744 | -9.2062
5.1581 | -34.1747 | -9.2167
5.1642 | -34.175 | -9.2272
5.1703 | -34.1752 | -9.2375
5.1764 | -34.1756 | -9.2477
5.1825 | -34.1758 | -9.2579
5.1886 | -34.1761 | -9.2679
5.1947 | -34.1764 | -9.2783
5.2008 | -34.1767 | -9.2888
5.2069 | -34.177 | -9.2993
5.213 | -34.1773 | -9.3095
5.2191 | -34.1776 | -9.3199
5.2252 | -34.1779 | -9.3304
5.2314 | -34.1782 | -9.341
5.2375 | -34.1785 | -9.3511
5.2436 | -34.1787 | -9.3615
5.2497 | -34.179 | -9.3715
5.2558 | -34.1794 | -9.3823
5.2619 | -34.1796 | -9.3929
5.268 | -34.18 | -9.4035
5.2741 | -34.1803 | -9.4139
5.2802 | -34.1806 | -9.4246
5.2863 | -34.1809 | -9.4349
5.2924 | -34.1812 | -9.4455
5.2985 | -34.1816 | -9.456
5.3046 | -34.1819 | -9.4664
5.3107 | -34.1821 | -9.477
5.3168 | -34.1825 | -9.4874
5.3229 | -34.1827 | -9.4979
5.329 | -34.183 | -9.5083
5.3351 | -34.1834 | -9.5184
5.3412 | -34.1836 | -9.5288
5.3473 | -34.1839 | -9.5392
5.3534 | -34.1843 | -9.5497
5.3595 | -34.1845 | -9.5602
5.3656 | -34.1849 | -9.5705
5.3717 | -34.1852 | -9.5812
5.3779 | -34.1855 | -9.592
5.384 | -34.1858 | -9.6024
5.3901 | -34.1861 | -9.613
5.3962 | -34.1864 | -9.6233
5.4023 | -34.1867 | -9.6339
5.4084 | -34.1871 | -9.6442
5.4145 | -34.1874 | -9.6544
5.4206 | -34.1877 | -9.6649
5.4267 | -34.188 | -9.6756
5.4328 | -34.1883 | -9.6861
5.4389 | -34.1886 | -9.6962
5.445 | -34.1889 | -9.7069
5.4511 | -34.1892 | -9.7172
5.4572 | -34.1895 | -9.7277
5.4633 | -34.1898 | -9.7381
5.4694 | -34.1901 | -9.7486
5.4755 | -34.1904 | -9.7588
5.4816 | -34.1907 | -9.7693
5.4877 | -34.191 | -9.7795
5.4938 | -34.1913 | -9.7897
5.4999 | -34.1916 | -9.8
5.506 | -34.192 | -9.8104
5.5121 | -34.1923 | -9.8207
5.5183 | -34.1926 | -9.8308
5.5244 | -34.1928 | -9.8414
5.5305 | -34.1932 | -9.8517
5.5366 | -34.1935 | -9.8618
5.5427 | -34.1937 | -9.8724
5.5488 | -34.194 | -9.8828
5.5549 | -34.1943 | -9.8932
5.561 | -34.1947 | -9.9038
5.5671 | -34.195 | -9.9141
5.5732 | -34.1953 | -9.9247
5.5793 | -34.1956 | -9.9353
5.5854 | -34.1959 | -9.9456
5.5915 | -34.1962 | -9.9563
5.5976 | -34.1965 | -9.9668
5.6037 | -34.1968 | -9.9773
5.6098 | -34.1971 | -9.9876
5.6159 | -34.1974 | -9.9978
5.622 | -34.1977 | -10.0083
5.6281 | -34.198 | -10.0187
5.6342 | -34.1983 | -10.0294
5.6403 | -34.1986 | -10.0399
5.6464 | -34.1989 | -10.0504
5.6525 | -34.1992 | -10.0609
5.6586 | -34.1995 | -10.071
5.6648 | -34.1999 | -10.0815
5.6709 | -34.2002 | -10.0921
5.677 | -34.2005 | -10.1025
5.6831 | -34.2008 | -10.1129
5.6892 | -34.2011 | -10.1235
5.6953 | -34.2014 | -10.1339
5.7014 | -34.2017 | -10.1443
5.7075 | -34.202 | -10.1548
5.7136 | -34.2023 | -10.1651
5.7197 | -34.2026 | -10.1755
5.7258 | -34.2028 | -10.1857
5.7319 | -34.2032 | -10.1961
5.738 | -34.2035 | -10.2065
5.7441 | -34.2038 | -10.2168
5.7502 | -34.2041 | -10.227
5.7563 | -34.2044 | -10.2374
5.7624 | -34.2048 | -10.248
5.7685 | -34.2051 | -10.2584
5.7746 | -34.2053 | -10.2691
5.7807 | -34.2056 | -10.2795
5.7868 | -34.2059 | -10.2901
5.7929 | -34.2062 | -10.3004
5.799 | -34.2065 | -10.3106
5.8052 | -34.2068 | -10.3209
5.8113 | -34.2071 | -10.3311
5.8174 | -34.2073 | -10.3414
5.8235 | -34.2077 | -10.3517
5.8296 | -34.2079 | -10.3623
5.8357 | -34.2082 | -10.3725
5.8418 | -34.2085 | -10.383
5.8479 | -34.2088 | -10.3935
5.854 | -34.2091 | -10.404
5.8601 | -34.2094 | -10.4143
5.8662 | -34.2097 | -10.4248
5.8723 | -34.21 | -10.435
5.8784 | -34.2103 | -10.4453
5.8845 | -34.2106 | -10.4556
5.8906 | -34.2108 | -10.466
5.8967 | -34.2111 | -10.4762
5.9028 | -34.2115 | -10.4864
5.9089 | -34.2118 | -10.4966
5.915 | -34.2121 | -10.5068
5.9211 | -34.2124 | -10.5171
5.9272 | -34.2127 | -10.5272
5.9333 | -34.213 | -10.5374
5.9394 | -34.2133 | -10.5477
5.9455 | -34.2136 | -10.5578
5.9517 | -34.214 | -10.568
5.9578 | -34.2143 | -10.5785
5.9639 | -34.2146 | -10.5885
5.97 | -34.215 | -10.5989
5.9761 | -34.2153 | -10.6091
5.9822 | -34.2156 | -10.6195
5.9883 | -34.2159 | -10.6299
5.9944 | -34.2162 | -10.6403
6.0005 | -34.2165 | -10.6507
6.0066 | -34.2168 | -10.6609
6.0127 | -34.2172 | -10.6709
6.0188 | -34.2175 | -10.6814
6.0249 | -34.2178 | -10.6918
6.031 | -34.2181 | -10.7019
6.0371 | -34.2184 | -10.7122
6.0432 | -34.2187 | -10.7224
6.0493 | -34.219 | -10.7326
6.0554 | -34.2193 | -10.7427
6.0615 | -34.2197 | -10.7531
6.0676 | -34.22 | -10.7635
6.0737 | -34.2203 | -10.7739
6.0798 | -34.2206 | -10.7842
6.0859 | -34.2209 | -10.7944
6.0921 | -34.2212 | -10.8045
6.0982 | -34.2215 | -10.8147
6.1043 | -34.2219 | -10.8249
6.1104 | -34.2222 | -10.8354
6.1165 | -34.2225 | -10.8453
6.1226 | -34.2229 | -10.8555
6.1287 | -34.2232 | -10.8657
6.1348 | -34.2235 | -10.8759
6.1409 | -34.2238 | -10.8861
6.147 | -34.2241 | -10.8963
6.1531 | -34.2244 | -10.9066
6.1592 | -34.2247 | -10.9167
6.1653 | -34.225 | -10.9268
6.1714 | -34.2253 | -10.9371
6.1775 | -34.2257 | -10.9475
6.1836 | -34.226 | -10.9577
6.1897 | -34.2263 | -10.968
6.1958 | -34.2266 | -10.9785
6.2019 | -34.227 | -10.9888
6.208 | -34.2273 | -10.999
6.2141 | -34.2276 | -11.0092
6.2202 | -34.228 | -11.0195
6.2263 | -34.2283 | -11.0297
6.2325 | -34.2286 | -11.0397
6.2386 | -34.2289 | -11.0497
6.2447 | -34.2292 | -11.0602
6.2508 | -34.2296 | -11.0704
6.2569 | -34.2299 | -11.0806
6.263 | -34.2302 | -11.0908
6.2691 | -34.2305 | -11.1009
6.2752 | -34.2308 | -11.1112
6.2813 | -34.2311 | -11.1214
6.2874 | -34.2314 | -11.1315
6.2935 | -34.2318 | -11.1417
6.2996 | -34.232 | -11.1519
6.3057 | -34.2324 | -11.1621
6.3118 | -34.2327 | -11.1725
6.3179 | -34.233 | -11.1827
6.324 | -34.2333 | -11.1929
6.3301 | -34.2336 | -11.2032
6.3362 | -34.2339 | -11.2133
6.3423 | -34.2343 | -11.2236
6.3484 | -34.2346 | -11.2339
6.3545 | -34.2349 | -11.2439
6.3606 | -34.2352 | -11.2538
6.3667 | -34.2355 | -11.2639
6.3728 | -34.2358 | -11.2741
6.379 | -34.2361 | -11.2844
6.3851 | -34.2365 | -11.2943
6.3912 | -34.2368 | -11.3045
6.3973 | -34.2371 | -11.3147
6.4034 | -34.2374 | -11.3248
6.4095 | -34.2377 | -11.3349
6.4156 | -34.238 | -11.3451
6.4217 | -34.2383 | -11.3553
6.4278 | -34.2386 | -11.3655
6.4339 | -34.239 | -11.3756
6.44 | -34.2393 | -11.3857
6.4461 | -34.2396 | -11.396
6.4522 | -34.2399 | -11.406
6.4583 | -34.2402 | -11.4163
6.4644 | -34.2405 | -11.4266
6.4705 | -34.2408 | -11.4365
6.4766 | -34.2411 | -11.4467
6.4827 | -34.2414 | -11.4569
6.4888 | -34.2417 | -11.4669
6.4949 | -34.242 | -11.4768
6.501 | -34.2424 | -11.4871
6.5071 | -34.2427 | -11.4971
6.5132 | -34.243 | -11.5073
6.5194 | -34.2433 | -11.5175
6.5255 | -34.2436 | -11.5274
6.5316 | -34.2439 | -11.5377
6.5377 | -34.2442 | -11.5479
6.5438 | -34.2445 | -11.5579
6.5499 | -34.2449 | -11.5679
6.556 | -34.2452 | -11.578
6.5621 | -34.2455 | -11.5884
6.5682 | -34.2458 | -11.5986
6.5743 | -34.2462 | -11.6087
6.5804 | -34.2465 | -11.6189
6.5865 | -34.2468 | -11.6291
6.5926 | -34.2471 | -11.6392
6.5987 | -34.2474 | -11.6491
6.6048 | -34.2477 | -11.6593
6.6109 | -34.2481 | -11.6695
6.617 | -34.2484 | -11.6798
6.6231 | -34.2486 | -11.6899
6.6292 | -34.249 | -11.7
6.6353 | -34.2493 | -11.7105
6.6414 | -34.2496 | -11.7205
6.6475 | -34.25 | -11.7307
6.6536 | -34.2502 | -11.7408
6.6597 | -34.2506 | -11.7507
6.6659 | -34.2509 | -11.7608
6.672 | -34.2512 | -11.7709
6.6781 | -34.2515 | -11.7811
6.6842 | -34.2518 | -11.7911
6.6903 | -34.2522 | -11.8011
6.6964 | -34.2525 | -11.8112
6.7025 | -34.2528 | -11.8212
6.7086 | -34.2531 | -11.8313
6.7147 | -34.2534 | -11.8415
6.7208 | -34.2537 | -11.8517
6.7269 | -34.2541 | -11.8617
6.733 | -34.2544 | -11.8717
6.7391 | -34.2547 | -11.8818
6.7452 | -34.255 | -11.8917
6.7513 | -34.2553 | -11.9016
6.7574 | -34.2557 | -11.9117
6.7635 | -34.256 | -11.9217
6.7696 | -34.2563 | -11.9317
6.7757 | -34.2567 | -11.9418
6.7818 | -34.257 | -11.9519
6.7879 | -34.2573 | -11.9618
6.794 | -34.2576 | -11.9718
6.8001 | -34.2579 | -11.9818
6.8063 | -34.2582 | -11.9919
6.8124 | -34.2585 | -12.0022
6.8185 | -34.2588 | -12.0124
6.8246 | -34.2591 | -12.0224
6.8307 | -34.2595 | -12.0324
6.8368 | -34.2598 | -12.0427
6.8429 | -34.2601 | -12.0528
6.849 | -34.2604 | -12.0629
6.8551 | -34.2608 | -12.0729
6.8612 | -34.2611 | -12.083
6.8673 | -34.2614 | -12.093
6.8734 | -34.2617 | -12.1033
6.8795 | -34.2621 | -12.1133
6.8856 | -34.2624 | -12.1234
6.8917 | -34.2627 | -12.1335
6.8978 | -34.263 | -12.1433
6.9039 | -34.2633 | -12.1531
6.91 | -34.2637 | -12.1633
6.9161 | -34.264 | -12.1731
6.9222 | -34.2643 | -12.1832
6.9283 | -34.2647 | -12.193
6.9344 | -34.265 | -12.203
6.9405 | -34.2654 | -12.213
6.9466 | -34.2657 | -12.2232
6.9528 | -34.266 | -12.2335
6.9589 | -34.2664 | -12.2436
6.965 | -34.2667 | -12.2538
6.9711 | -34.2671 | -12.2639
6.9772 | -34.2674 | -12.2739
6.9833 | -34.2677 | -12.2839
6.9894 | -34.2681 | -12.2942
6.9955 | -34.2684 | -12.304
7.0016 | -34.2687 | -12.314
7.0077 | -34.2691 | -12.324
7.0138 | -34.2694 | -12.334
7.0199 | -34.2697 | -12.344
7.026 | -34.2701 | -12.354
7.0321 | -34.2704 | -12.3638
7.0382 | -34.2707 | -12.3736
7.0443 | -34.2711 | -12.3837
7.0504 | -34.2714 | -12.3938
7.0565 | -34.2717 | -12.4038
7.0626 | -34.2721 | -12.4136
7.0687 | -34.2724 | -12.4233
7.0748 | -34.2728 | -12.433
7.0809 | -34.2731 | -12.443
7.087 | -34.2734 | -12.4531
7.0932 | -34.2737 | -12.4631
7.0993 | -34.274 | -12.4731
7.1054 | -34.2744 | -12.483
7.1115 | -34.2748 | -12.4928
7.1176 | -34.2751 | -12.5027
7.1237 | -34.2754 | -12.5126
7.1298 | -34.2757 | -12.5225
7.1359 | -34.2761 | -12.5325
7.142 | -34.2764 | -12.5422
7.1481 | -34.2767 | -12.552
7.1542 | -34.277 | -12.562
7.1603 | -34.2774 | -12.572
7.1664 | -34.2777 | -12.5819
7.1725 | -34.2781 | -12.5918
7.1786 | -34.2784 | -12.6017
7.1847 | -34.2787 | -12.6118
7.1908 | -34.2791 | -12.6219
7.1969 | -34.2794 | -12.632
7.203 | -34.2797 | -12.6417
7.2091 | -34.2801 | -12.6515
7.2152 | -34.2804 | -12.6614
7.2213 | -34.2808 | -12.6712
7.2274 | -34.2811 | -12.6811
7.2335 | -34.2814 | -12.691
7.2397 | -34.2818 | -12.7008
7.2458 | -34.2822 | -12.7105
7.2519 | -34.2825 | -12.7204
7.258 | -34.2828 | -12.7305
7.2641 | -34.2832 | -12.7404
7.2702 | -34.2835 | -12.7505
7.2763 | -34.2838 | -12.7604
7.2824 | -34.2842 | -12.7704
7.2885 | -34.2845 | -12.7803
7.2946 | -34.2849 | -12.7902
7.3007 | -34.2852 | -12.8002
7.3068 | -34.2855 | -12.8103
7.3129 | -34.2859 | -12.8203
7.319 | -34.2862 | -12.8302
7.3251 | -34.2866 | -12.8399
7.3312 | -34.2869 | -12.8498
7.3373 | -34.2873 | -12.8595
7.3434 | -34.2876 | -12.8693
7.3495 | -34.288 | -12.879
7.3556 | -34.2883 | -12.889
7.3617 | -34.2887 | -12.8991
7.3678 | -34.289 | -12.9089
7.3739 | -34.2893 | -12.9189
7.3801 | -34.2897 | -12.9291
7.3862 | -34.29 | -12.9389
7.3923 | -34.2904 | -12.9489
7.3984 | -34.2907 | -12.9586
7.4045 | -34.2911 | -12.9685
7.4106 | -34.2915 | -12.9783
7.4167 | -34.2918 | -12.9883
7.4228 | -34.2922 | -12.9982
7.4289 | -34.2925 | -13.0081
7.435 | -34.2929 | -13.0178
7.4411 | -34.2932 | -13.0277
7.4472 | -34.2935 | -13.0376
7.4533 | -34.2939 | -13.0472
7.4594 | -34.2943 | -13.0573
7.4655 | -34.2946 | -13.067
7.4716 | -34.295 | -13.0771
7.4777 | -34.2953 | -13.0869
7.4838 | -34.2956 | -13.0966
7.4899 | -34.296 | -13.1064
7.496 | -34.2963 | -13.116
7.5021 | -34.2967 | -13.1259
7.5082 | -34.297 | -13.1357
7.5143 | -34.2974 | -13.1456
7.5204 | -34.2977 | -13.1556
7.5266 | -34.298 | -13.1653
7.5327 | -34.2984 | -13.175
7.5388 | -34.2987 | -13.1851
7.5449 | -34.2991 | -13.1952
7.551 | -34.2994 | -13.2049
7.5571 | -34.2998 | -13.2148
7.5632 | -34.3002 | -13.2247
7.5693 | -34.3005 | -13.2344
7.5754 | -34.3009 | -13.2441
7.5815 | -34.3012 | -13.254
7.5876 | -34.3016 | -13.2641
7.5937 | -34.3019 | -13.274
7.5998 | -34.3023 | -13.2838
7.6059 | -34.3027 | -13.2938
7.612 | -34.303 | -13.3037
7.6181 | -34.3034 | -13.3134
7.6242 | -34.3037 | -13.3236
7.6303 | -34.3041 | -13.3335
7.6364 | -34.3044 | -13.3435
7.6425 | -34.3048 | -13.3534
7.6486 | -34.3051 | -13.3633
7.6547 | -34.3055 | -13.373
7.6608 | -34.3058 | -13.3829
7.667 | -34.3061 | -13.3928
7.6731 | -34.3065 | -13.4027
7.6792 | -34.3069 | -13.4128
7.6853 | -34.3072 | -13.4228
7.6914 | -34.3076 | -13.4327
7.6975 | -34.3079 | -13.4426
7.7036 | -34.3083 | -13.4525
7.7097 | -34.3087 | -13.4626
7.7158 | -34.309 | -13.4725
7.7219 | -34.3094 | -13.4822
7.728 | -34.3097 | -13.4921
7.7341 | -34.3101 | -13.502
7.7402 | -34.3105 | -13.512
7.7463 | -34.3108 | -13.5219
7.7524 | -34.3112 | -13.532
7.7585 | -34.3115 | -13.5419
7.7646 | -34.3119 | -13.5518
7.7707 | -34.3122 | -13.5618
7.7768 | -34.3125 | -13.5718
7.7829 | -34.3129 | -13.5819
7.789 | -34.3132 | -13.592
7.7951 | -34.3136 | -13.6019
7.8012 | -34.3139 | -13.6118
7.8073 | -34.3143 | -13.6219
7.8135 | -34.3147 | -13.6318
7.8196 | -34.3151 | -13.6418
7.8257 | -34.3154 | -13.6519
7.8318 | -34.3157 | -13.6619
7.8379 | -34.3161 | -13.6717
7.844 | -34.3165 | -13.6816
7.8501 | -34.3168 | -13.6917
7.8562 | -34.3172 | -13.7018
7.8623 | -34.3176 | -13.7116
7.8684 | -34.3179 | -13.7216
7.8745 | -34.3182 | -13.7316
7.8806 | -34.3186 | -13.7416
7.8867 | -34.319 | -13.7514
7.8928 | -34.3193 | -13.7611
7.8989 | -34.3197 | -13.771
7.905 | -34.32 | -13.781
7.9111 | -34.3204 | -13.791
7.9172 | -34.3207 | -13.8009
7.9233 | -34.3211 | -13.8109
7.9294 | -34.3214 | -13.821
7.9355 | -34.3218 | -13.8307
7.9416 | -34.3221 | -13.8407
7.9477 | -34.3225 | -13.8505
7.9539 | -34.3228 | -13.8603
7.96 | -34.3232 | -13.8703
7.9661 | -34.3235 | -13.8802
7.9722 | -34.3239 | -13.89
7.9783 | -34.3243 | -13.9001
7.9844 | -34.3247 | -13.9099
7.9905 | -34.325 | -13.9201
7.9966 | -34.3253 | -13.9301
8.0027 | -34.3257 | -13.94
8.0088 | -34.326 | -13.9499
8.0149 | -34.3264 | -13.9599
8.021 | -34.3267 | -13.9699
8.0271 | -34.3271 | -13.98
8.0332 | -34.3275 | -13.9898
8.0393 | -34.3278 | -13.9999
8.0454 | -34.3282 | -14.01
8.0515 | -34.3286 | -14.0199
8.0576 | -34.3289 | -14.0301
8.0637 | -34.3293 | -14.0399
8.0698 | -34.3296 | -14.0497
8.0759 | -34.33 | -14.0596
8.082 | -34.3303 | -14.0695
8.0881 | -34.3307 | -14.0794
8.0942 | -34.3311 | -14.0894
8.1004 | -34.3315 | -14.0994
8.1065 | -34.3318 | -14.1093
8.1126 | -34.3321 | -14.1193
8.1187 | -34.3325 | -14.1293
8.1248 | -34.3328 | -14.139
8.1309 | -34.3332 | -14.1488
8.137 | -34.3335 | -14.1586
8.1431 | -34.3339 | -14.1687
8.1492 | -34.3343 | -14.1786
8.1553 | -34.3346 | -14.1886
8.1614 | -34.335 | -14.1987
8.1675 | -34.3353 | -14.2086
8.1736 | -34.3357 | -14.2187
8.1797 | -34.336 | -14.2287
8.1858 | -34.3364 | -14.2387
8.1919 | -34.3367 | -14.2486
8.198 | -34.3371 | -14.2584
8.2041 | -34.3375 | -14.2684
8.2102 | -34.3378 | -14.2783
8.2163 | -34.3382 | -14.2881
8.2224 | -34.3385 | -14.2981
8.2285 | -34.3389 | -14.3078
8.2346 | -34.3392 | -14.3177
8.2408 | -34.3396 | -14.3275
8.2469 | -34.34 | -14.3375
8.253 | -34.3403 | -14.3473
8.2591 | -34.3407 | -14.3573
8.2652 | -34.3411 | -14.3673
8.2713 | -34.3414 | -14.3772
8.2774 | -34.3418 | -14.3873
8.2835 | -34.3422 | -14.3974
8.2896 | -34.3425 | -14.4071
8.2957 | -34.3429 | -14.4172
8.3018 | -34.3432 | -14.4271
8.3079 | -34.3436 | -14.4369
8.314 | -34.344 | -14.4469
8.3201 | -34.3444 | -14.4567
8.3262 | -34.3447 | -14.4667
8.3323 | -34.3451 | -14.4766
8.3384 | -34.3454 | -14.4866
8.3445 | -34.3458 | -14.4964
8.3506 | -34.3462 | -14.5064
8.3567 | -34.3465 | -14.5164
8.3628 | -34.3469 | -14.5263
8.3689 | -34.3472 | -14.5361
8.375 | -34.3476 | -14.5461
8.3812 | -34.3479 | -14.5558
8.3873 | -34.3483 | -14.5657
8.3934 | -34.3487 | -14.5757
8.3995 | -34.349 | -14.5855
8.4056 | -34.3494 | -14.5954
8.4117 | -34.3497 | -14.6053
8.4178 | -34.3501 | -14.6153
8.4239 | -34.3504 | -14.625
8.43 | -34.3508 | -14.635
8.4361 | -34.3511 | -14.645
8.4422 | -34.3515 | -14.6548
8.4483 | -34.3519 | -14.6646
8.4544 | -34.3522 | -14.6745
8.4605 | -34.3526 | -14.6843
8.4666 | -34.3529 | -14.6944
8.4727 | -34.3533 | -14.7043
8.4788 | -34.3537 | -14.7142
8.4849 | -34.354 | -14.7241
8.491 | -34.3544 | -14.7339
8.4971 | -34.3548 | -14.7437
8.5032 | -34.3551 | -14.7535
8.5093 | -34.3555 | -14.7632
8.5154 | -34.3559 | -14.7732
8.5215 | -34.3562 | -14.7829
8.5277 | -34.3566 | -14.7928
8.5338 | -34.357 | -14.8026
8.5399 | -34.3573 | -14.8125
8.546 | -34.3577 | -14.8222
8.5521 | -34.3581 | -14.8321
8.5582 | -34.3584 | -14.842
8.5643 | -34.3588 | -14.8516
8.5704 | -34.3591 | -14.8613
8.5765 | -34.3595 | -14.8712
8.5826 | -34.3599 | -14.8808
8.5887 | -34.3603 | -14.8906
8.5948 | -34.3606 | -14.9005
8.6009 | -34.361 | -14.9101
8.607 | -34.3614 | -14.92
8.6131 | -34.3617 | -14.9299
8.6192 | -34.3621 | -14.9396
8.6253 | -34.3625 | -14.9494
8.6314 | -34.3629 | -14.9592
8.6375 | -34.3632 | -14.969
8.6436 | -34.3636 | -14.9789
8.6497 | -34.364 | -14.9886
8.6558 | -34.3643 | -14.9984
8.6619 | -34.3646 | -15.0083
8.6681 | -34.365 | -15.0183
8.6742 | -34.3654 | -15.0281
8.6803 | -34.3658 | -15.0378
8.6864 | -34.3662 | -15.0474
8.6925 | -34.3665 | -15.057
8.6986 | -34.3669 | -15.0669
8.7047 | -34.3673 | -15.0768
8.7108 | -34.3676 | -15.0866
8.7169 | -34.368 | -15.0965
8.723 | -34.3684 | -15.1062
8.7291 | -34.3687 | -15.1158
8.7352 | -34.3691 | -15.1258
8.7413 | -34.3694 | -15.1357
8.7474 | -34.3698 | -15.1453
8.7535 | -34.3702 | -15.1551
8.7596 | -34.3706 | -15.165
8.7657 | -34.3709 | -15.1749
8.7718 | -34.3713 | -15.1848
8.7779 | -34.3717 | -15.1946
8.784 | -34.3721 | -15.2044
8.7901 | -34.3725 | -15.2142
8.7962 | -34.3728 | -15.2238
8.8023 | -34.3732 | -15.2335
8.8084 | -34.3736 | -15.2432
8.8146 | -34.3739 | -15.253
8.8207 | -34.3743 | -15.2628
8.8268 | -34.3747 | -15.2723
8.8329 | -34.375 | -15.282
8.839 | -34.3754 | -15.2918
8.8451 | -34.3758 | -15.3015
8.8512 | -34.3762 | -15.3111
8.8573 | -34.3765 | -15.3209
8.8634 | -34.3769 | -15.3305
8.8695 | -34.3773 | -15.3404
8.8756 | -34.3777 | -15.3502
8.8817 | -34.378 | -15.3599
8.8878 | -34.3784 | -15.3698
8.8939 | -34.3788 | -15.3795
8.9 | -34.3792 | -15.3893
8.9061 | -34.3795 | -15.399
8.9122 | -34.3799 | -15.4087
8.9183 | -34.3803 | -15.4185
8.9244 | -34.3807 | -15.4284
8.9305 | -34.3811 | -15.4379
8.9366 | -34.3815 | -15.4477
8.9427 | -34.3819 | -15.4575
8.9488 | -34.3823 | -15.4672
8.955 | -34.3826 | -15.4769
8.9611 | -34.383 | -15.4868
8.9672 | -34.3834 | -15.4965
8.9733 | -34.3838 | -15.5061
8.9794 | -34.3842 | -15.5157
8.9855 | -34.3846 | -15.5255
8.9916 | -34.385 | -15.5354
8.9977 | -34.3854 | -15.545
9.0038 | -34.3858 | -15.5546
9.0099 | -34.3862 | -15.5644
9.016 | -34.3865 | -15.5741
9.0221 | -34.3869 | -15.5837
9.0282 | -34.3873 | -15.5932
9.0343 | -34.3877 | -15.6029
9.0404 | -34.3881 | -15.6126
9.0465 | -34.3885 | -15.6225
9.0526 | -34.3889 | -15.6321
9.0587 | -34.3893 | -15.6419
9.0648 | -34.3897 | -15.6514
9.0709 | -34.3901 | -15.6611
9.077 | -34.3905 | -15.6709
9.0831 | -34.3909 | -15.6807
9.0892 | -34.3913 | -15.6904
9.0953 | -34.3917 | -15.7002
9.1015 | -34.3921 | -15.7099
9.1076 | -34.3925 | -15.7194
9.1137 | -34.3929 | -15.7291
9.1198 | -34.3933 | -15.7387
9.1259 | -34.3937 | -15.7484
9.132 | -34.3941 | -15.7582
9.1381 | -34.3945 | -15.7678
9.1442 | -34.3949 | -15.7773
9.1503 | -34.3953 | -15.7871
9.1564 | -34.3957 | -15.7967
9.1625 | -34.3961 | -15.8065
9.1686 | -34.3965 | -15.8161
9.1747 | -34.3969 | -15.8258
9.1808 | -34.3973 | -15.8354
9.1869 | -34.3977 | -15.8451
9.193 | -34.3981 | -15.8548
9.1991 | -34.3985 | -15.8646
9.2052 | -34.3989 | -15.8742
9.2113 | -34.3993 | -15.8839
9.2174 | -34.3998 | -15.8936
9.2235 | -34.4002 | -15.9032
9.2296 | -34.4006 | -15.9129
9.2357 | -34.401 | -15.9227
9.2419 | -34.4014 | -15.9324
9.248 | -34.4018 | -15.942
9.2541 | -34.4022 | -15.9517
9.2602 | -34.4026 | -15.9615
9.2663 | -34.403 | -15.971
9.2724 | -34.4035 | -15.9806
9.2785 | -34.4038 | -15.9902
9.2846 | -34.4043 | -15.9998
9.2907 | -34.4047 | -16.0094
9.2968 | -34.4051 | -16.0192
9.3029 | -34.4055 | -16.029
9.309 | -34.4059 | -16.0386
9.3151 | -34.4063 | -16.0481
9.3212 | -34.4067 | -16.0577
9.3273 | -34.4072 | -16.0673
9.3334 | -34.4076 | -16.0768
9.3395 | -34.408 | -16.0863
9.3456 | -34.4084 | -16.0958
9.3517 | -34.4088 | -16.1056
9.3578 | -34.4092 | -16.1152
9.3639 | -34.4096 | -16.1247
9.37 | -34.41 | -16.1343
9.3761 | -34.4104 | -16.1441
9.3822 | -34.4109 | -16.1538
9.3884 | -34.4113 | -16.1632
9.3945 | -34.4117 | -16.1729
9.4006 | -34.4121 | -16.1826
9.4067 | -34.4125 | -16.1922
9.4128 | -34.4129 | -16.2019
9.4189 | -34.4134 | -16.2117
9.425 | -34.4138 | -16.2212
9.4311 | -34.4142 | -16.2307
9.4372 | -34.4147 | -16.2403
9.4433 | -34.4151 | -16.2499
9.4494 | -34.4155 | -16.2594
9.4555 | -34.4159 | -16.2691
9.4616 | -34.4163 | -16.2787
9.4677 | -34.4168 | -16.2883
9.4738 | -34.4172 | -16.2979
9.4799 | -34.4176 | -16.3076
9.486 | -34.418 | -16.3172
9.4921 | -34.4184 | -16.3268
9.4982 | -34.4189 | -16.3362
9.5043 | -34.4193 | -16.3457
9.5104 | -34.4197 | -16.3552
9.5165 | -34.4201 | -16.3648
9.5226 | -34.4205 | -16.3745
9.5288 | -34.4209 | -16.384
9.5349 | -34.4214 | -16.3937
9.541 | -34.4218 | -16.4031
9.5471 | -34.4222 | -16.4126
9.5532 | -34.4226 | -16.4223
9.5593 | -34.423 | -16.4321
9.5654 | -34.4235 | -16.4416
9.5715 | -34.4239 | -16.4511
9.5776 | -34.4243 | -16.4609
9.5837 | -34.4247 | -16.4704
9.5898 | -34.4251 | -16.4801
9.5959 | -34.4255 | -16.4897
9.602 | -34.4259 | -16.4991
9.6081 | -34.4264 | -16.5088
9.6142 | -34.4268 | -16.5183
9.6203 | -34.4272 | -16.5279
9.6264 | -34.4276 | -16.5374
9.6325 | -34.428 | -16.5471
9.6386 | -34.4284 | -16.5565
9.6447 | -34.4288 | -16.5662
9.6508 | -34.4293 | -16.5756
9.6569 | -34.4297 | -16.5852
9.663 | -34.4301 | -16.5948
9.6691 | -34.4306 | -16.6044
9.6753 | -34.431 | -16.6139
9.6814 | -34.4314 | -16.6234
9.6875 | -34.4318 | -16.6327
9.6936 | -34.4323 | -16.6421
9.6997 | -34.4327 | -16.6518
9.7058 | -34.4331 | -16.6613
9.7119 | -34.4336 | -16.671
9.718 | -34.434 | -16.6807
9.7241 | -34.4344 | -16.6901
9.7302 | -34.4348 | -16.6997
9.7363 | -34.4353 | -16.7092
9.7424 | -34.4357 | -16.7186
9.7485 | -34.4361 | -16.7282
9.7546 | -34.4365 | -16.7375
9.7607 | -34.437 | -16.7473
9.7668 | -34.4374 | -16.7569
9.7729 | -34.4378 | -16.7663
9.779 | -34.4383 | -16.7757
9.7851 | -34.4387 | -16.7854
9.7912 | -34.4391 | -16.7949
9.7973 | -34.4396 | -16.8045
9.8034 | -34.44 | -16.814
9.8095 | -34.4404 | -16.8237
9.8157 | -34.4409 | -16.8332
9.8218 | -34.4413 | -16.843
9.8279 | -34.4417 | -16.8523
9.834 | -34.4421 | -16.8619
9.8401 | -34.4426 | -16.8715
9.8462 | -34.443 | -16.8808
9.8523 | -34.4435 | -16.8903
9.8584 | -34.4439 | -16.8998
9.8645 | -34.4443 | -16.9095
9.8706 | -34.4447 | -16.9191
9.8767 | -34.4451 | -16.9285
9.8828 | -34.4455 | -16.938
9.8889 | -34.446 | -16.9473
9.895 | -34.4464 | -16.9569
9.9011 | -34.4468 | -16.9664
9.9072 | -34.4472 | -16.9759
9.9133 | -34.4476 | -16.9854
9.9194 | -34.4481 | -16.9947
9.9255 | -34.4485 | -17.0042
9.9316 | -34.4489 | -17.0137
9.9377 | -34.4493 | -17.0232
9.9438 | -34.4497 | -17.0328
9.9499 | -34.4502 | -17.0425
9.956 | -34.4506 | -17.0521
9.9622 | -34.451 | -17.0617
9.9683 | -34.4514 | -17.0712
9.9744 | -34.4518 | -17.0808
9.9805 | -34.4523 | -17.0903
9.9866 | -34.4527 | -17.0997
9.9927 | -34.4531 | -17.1091
9.9988 | -34.4535 | -17.1189
10.0049 | -34.4539 | -17.1284
10.011 | -34.4543 | -17.1379
10.0171 | -34.4548 | -17.1474
10.0232 | -34.4552 | -17.1568
10.0293 | -34.4557 | -17.1664
10.0354 | -34.4561 | -17.1759
10.0415 | -34.4566 | -17.1854
10.0476 | -34.457 | -17.195
10.0537 | -34.4574 | -17.2046
10.0598 | -34.4578 | -17.214
10.0659 | -34.4582 | -17.2235
10.072 | -34.4586 | -17.233
10.0781 | -34.459 | -17.2424
10.0842 | -34.4595 | -17.2519
10.0903 | -34.4599 | -17.2613
10.0964 | -34.4603 | -17.2706
10.1026 | -34.4607 | -17.2802
10.1087 | -34.4612 | -17.2896
10.1148 | -34.4616 | -17.2992
10.1209 | -34.462 | -17.3085
10.127 | -34.4624 | -17.318
10.1331 | -34.4628 | -17.3273
10.1392 | -34.4632 | -17.3367
10.1453 | -34.4637 | -17.3462
10.1514 | -34.4641 | -17.3556
10.1575 | -34.4645 | -17.3652
10.1636 | -34.465 | -17.3745
10.1697 | -34.4654 | -17.3841
10.1758 | -34.4658 | -17.3935
10.1819 | -34.4663 | -17.4028
10.188 | -34.4667 | -17.4123
10.1941 | -34.4671 | -17.4217
10.2002 | -34.4675 | -17.4311
10.2063 | -34.4679 | -17.4406
10.2124 | -34.4683 | -17.4501
10.2185 | -34.4687 | -17.4596
10.2246 | -34.4691 | -17.469
10.2307 | -34.4696 | -17.4785
10.2368 | -34.47 | -17.4876
10.2429 | -34.4704 | -17.497
10.2491 | -34.4708 | -17.5065
10.2552 | -34.4712 | -17.516
10.2613 | -34.4716 | -17.5255
10.2674 | -34.4721 | -17.5348
10.2735 | -34.4725 | -17.5444
10.2796 | -34.4729 | -17.5538
10.2857 | -34.4733 | -17.5632
10.2918 | -34.4738 | -17.5728
10.2979 | -34.4742 | -17.5823
10.304 | -34.4746 | -17.5915
10.3101 | -34.475 | -17.601
10.3162 | -34.4755 | -17.6105
10.3223 | -34.4759 | -17.6199
10.3284 | -34.4763 | -17.6292
10.3345 | -34.4768 | -17.6386
10.3406 | -34.4772 | -17.6478
10.3467 | -34.4776 | -17.6572
10.3528 | -34.478 | -17.6666
10.3589 | -34.4784 | -17.6761
10.365 | -34.4789 | -17.6856
10.3711 | -34.4793 | -17.6951
10.3772 | -34.4797 | -17.7044
10.3833 | -34.4802 | -17.7138
10.3895 | -34.4806 | -17.7232
10.3956 | -34.481 | -17.7326
10.4017 | -34.4814 | -17.7418
10.4078 | -34.4818 | -17.7511
10.4139 | -34.4823 | -17.7605
10.42 | -34.4827 | -17.7701
10.4261 | -34.4831 | -17.7795
10.4322 | -34.4835 | -17.7888
10.4383 | -34.4839 | -17.7982
10.4444 | -34.4843 | -17.8076
10.4505 | -34.4847 | -17.817
10.4566 | -34.4851 | -17.8263
10.4627 | -34.4855 | -17.8356
10.4688 | -34.486 | -17.8449
10.4749 | -34.4864 | -17.8543
10.481 | -34.4868 | -17.8636
10.4871 | -34.4872 | -17.8729
10.4932 | -34.4877 | -17.8822
10.4993 | -34.4881 | -17.8916
10.5054 | -34.4885 | -17.9011
10.5115 | -34.4889 | -17.9104
10.5176 | -34.4894 | -17.9197
10.5237 | -34.4898 | -17.9291
10.5298 | -34.4902 | -17.9384
10.536 | -34.4906 | -17.9479
10.5421 | -34.491 | -17.9571
10.5482 | -34.4914 | -17.9666
10.5543 | -34.4919 | -17.9758
10.5604 | -34.4923 | -17.985
10.5665 | -34.4927 | -17.9943
10.5726 | -34.4931 | -18.0037
10.5787 | -34.4936 | -18.0131
10.5848 | -34.494 | -18.0223
10.5909 | -34.4944 | -18.0316
10.597 | -34.4948 | -18.041
10.6031 | -34.4952 | -18.0505
10.6092 | -34.4956 | -18.0599
10.6153 | -34.4961 | -18.0694
10.6214 | -34.4965 | -18.0783
10.6275 | -34.4969 | -18.0877
10.6336 | -34.4974 | -18.097
10.6397 | -34.4978 | -18.1064
10.6458 | -34.4982 | -18.1157
10.6519 | -34.4986 | -18.1252
10.658 | -34.499 | -18.1344
10.6641 | -34.4994 | -18.1439
10.6702 | -34.4999 | -18.1533
10.6764 | -34.5003 | -18.1627
10.6825 | -34.5007 | -18.1719
10.6886 | -34.5011 | -18.181
10.6947 | -34.5015 | -18.1905
10.7008 | -34.502 | -18.1999
10.7069 | -34.5024 | -18.2093
10.713 | -34.5028 | -18.2186
10.7191 | -34.5032 | -18.2279
10.7252 | -34.5037 | -18.237
10.7313 | -34.5041 | -18.2462
10.7374 | -34.5045 | -18.2556
10.7435 | -34.5049 | -18.2649
10.7496 | -34.5053 | -18.2742
10.7557 | -34.5057 | -18.2836
10.7618 | -34.5061 | -18.2928
10.7679 | -34.5066 | -18.3021
10.774 | -34.507 | -18.3116
10.7801 | -34.5074 | -18.3209
10.7862 | -34.5078 | -18.3301
10.7923 | -34.5083 | -18.3394
10.7984 | -34.5087 | -18.3487
10.8045 | -34.5091 | -18.3578
10.8106 | -34.5095 | -18.3672
10.8168 | -34.5099 | -18.3764
10.8229 | -34.5104 | -18.3858
10.829 | -34.5108 | -18.3951
10.8351 | -34.5112 | -18.4041
10.8412 | -34.5116 | -18.4133
10.8473 | -34.5121 | -18.4225
10.8534 | -34.5125 | -18.4317
10.8595 | -34.5129 | -18.4411
10.8656 | -34.5133 | -18.45
10.8717 | -34.5138 | -18.4593
10.8778 | -34.5142 | -18.4687
10.8839 | -34.5146 | -18.4779
10.89 | -34.515 | -18.4871
10.8961 | -34.5155 | -18.4964
10.9022 | -34.5159 | -18.5056
10.9083 | -34.5163 | -18.5148
10.9144 | -34.5167 | -18.524
10.9205 | -34.5171 | -18.5333
10.9266 | -34.5175 | -18.5425
10.9327 | -34.518 | -18.5517
10.9388 | -34.5184 | -18.5609
10.9449 | -34.5188 | -18.5701
10.951 | -34.5193 | -18.5794
10.9571 | -34.5197 | -18.5886
10.9633 | -34.5201 | -18.598
10.9694 | -34.5205 | -18.6073
10.9755 | -34.521 | -18.6164
10.9816 | -34.5214 | -18.6258
10.9877 | -34.5218 | -18.6348
10.9938 | -34.5222 | -18.6442
10.9999 | -34.5227 | -18.6534
11.006 | -34.5231 | -18.6623
11.0121 | -34.5235 | -18.6716
11.0182 | -34.524 | -18.6808
11.0243 | -34.5244 | -18.6899
11.0304 | -34.5248 | -18.6991
11.0365 | -34.5252 | -18.7082
11.0426 | -34.5257 | -18.7174
11.0487 | -34.5261 | -18.7267
11.0548 | -34.5266 | -18.7358
11.0609 | -34.527 | -18.7451
11.067 | -34.5274 | -18.7542
11.0731 | -34.5279 | -18.7634
11.0792 | -34.5283 | -18.7725
11.0853 | -34.5287 | -18.7817
11.0914 | -34.5291 | -18.791
11.0975 | -34.5296 | -18.8001
11.1037 | -34.53 | -18.8093
11.1098 | -34.5304 | -18.8186
11.1159 | -34.5309 | -18.8276
11.122 | -34.5313 | -18.8367
11.1281 | -34.5317 | -18.846
11.1342 | -34.5322 | -18.8552
11.1403 | -34.5326 | -18.8644
11.1464 | -34.533 | -18.8737
11.1525 | -34.5334 | -18.8826
11.1586 | -34.5339 | -18.8919
11.1647 | -34.5343 | -18.9009
11.1708 | -34.5347 | -18.9101
11.1769 | -34.5351 | -18.9193
11.183 | -34.5356 | -18.9282
11.1891 | -34.536 | -18.9373
11.1952 | -34.5365 | -18.9465
11.2013 | -34.5369 | -18.9557
11.2074 | -34.5373 | -18.9649
11.2135 | -34.5377 | -18.9742
11.2196 | -34.5382 | -18.9832
11.2257 | -34.5386 | -18.9924
11.2318 | -34.5391 | -19.0014
11.2379 | -34.5395 | -19.0107
11.244 | -34.54 | -19.0199
11.2502 | -34.5404 | -19.029
11.2563 | -34.5409 | -19.0381
11.2624 | -34.5413 | -19.0471
11.2685 | -34.5417 | -19.0563
11.2746 | -34.5422 | -19.0654
11.2807 | -34.5426 | -19.0747
11.2868 | -34.5431 | -19.0838
11.2929 | -34.5435 | -19.0927
11.299 | -34.5439 | -19.1019
11.3051 | -34.5444 | -19.1112
11.3112 | -34.5448 | -19.1204
11.3173 | -34.5453 | -19.1293
11.3234 | -34.5457 | -19.1381
11.3295 | -34.5461 | -19.1474
11.3356 | -34.5466 | -19.1567
11.3417 | -34.547 | -19.1656
11.3478 | -34.5474 | -19.1749
11.3539 | -34.5478 | -19.1842
11.36 | -34.5483 | -19.1932
11.3661 | -34.5487 | -19.2023
11.3722 | -34.5492 | -19.2112
11.3783 | -34.5496 | -19.2204
11.3844 | -34.55 | -19.2295
11.3906 | -34.5505 | -19.2387
11.3967 | -34.5509 | -19.2477
11.4028 | -34.5514 | -19.2569
11.4089 | -34.5518 | -19.2659
11.415 | -34.5522 | -19.275
11.4211 | -34.5526 | -19.2842
11.4272 | -34.5531 | -19.2933
11.4333 | -34.5535 | -19.3022
11.4394 | -34.5539 | -19.3113
11.4455 | -34.5544 | -19.3203
11.4516 | -34.5548 | -19.3294
11.4577 | -34.5552 | -19.3385
11.4638 | -34.5557 | -19.3473
11.4699 | -34.5561 | -19.3565
11.476 | -34.5566 | -19.3657
11.4821 | -34.557 | -19.3748
11.4882 | -34.5574 | -19.384
11.4943 | -34.5579 | -19.3931
11.5004 | -34.5583 | -19.402
11.5065 | -34.5588 | -19.411
11.5126 | -34.5592 | -19.4199
11.5187 | -34.5597 | -19.429
11.5248 | -34.5601 | -19.4382
11.5309 | -34.5606 | -19.4473
11.5371 | -34.561 | -19.4562
11.5432 | -34.5615 | -19.4653
11.5493 | -34.5619 | -19.4746
11.5554 | -34.5623 | -19.4836
11.5615 | -34.5628 | -19.4927
11.5676 | -34.5632 | -19.5019
11.5737 | -34.5637 | -19.5107
11.5798 | -34.5641 | -19.5196
11.5859 | -34.5646 | -19.5287
11.592 | -34.565 | -19.5377
11.5981 | -34.5654 | -19.5466
11.6042 | -34.5659 | -19.5557
11.6103 | -34.5663 | -19.5646
11.6164 | -34.5668 | -19.5735
11.6225 | -34.5672 | -19.5825
11.6286 | -34.5676 | -19.5915
11.6347 | -34.5681 | -19.6006
11.6408 | -34.5685 | -19.6095
11.6469 | -34.569 | -19.6185
11.653 | -34.5694 | -19.6275
11.6591 | -34.5698 | -19.6365
11.6652 | -34.5703 | -19.6455
11.6713 | -34.5707 | -19.6544
11.6775 | -34.5712 | -19.6632
11.6836 | -34.5716 | -19.6722
11.6897 | -34.5721 | -19.6811
11.6958 | -34.5725 | -19.6902
11.7019 | -34.5729 | -19.6994
11.708 | -34.5734 | -19.7083
11.7141 | -34.5738 | -19.7172
11.7202 | -34.5743 | -19.7263
11.7263 | -34.5747 | -19.7353
11.7324 | -34.5751 | -19.7444
11.7385 | -34.5756 | -19.7535
11.7446 | -34.576 | -19.7624
11.7507 | -34.5765 | -19.7711
11.7568 | -34.5769 | -19.7801
11.7629 | -34.5773 | -19.7891
11.769 | -34.5778 | -19.798
11.7751 | -34.5782 | -19.807
11.7812 | -34.5787 | -19.8157
11.7873 | -34.5791 | -19.8248
11.7934 | -34.5796 | -19.8338
11.7995 | -34.58 | -19.8427
11.8056 | -34.5805 | -19.8516
11.8117 | -34.5809 | -19.8606
11.8178 | -34.5814 | -19.8695
11.824 | -34.5818 | -19.8783
11.8301 | -34.5823 | -19.8872
11.8362 | -34.5827 | -19.896
11.8423 | -34.5831 | -19.905
11.8484 | -34.5836 | -19.9139
11.8545 | -34.5841 | -19.9228
11.8606 | -34.5845 | -19.9317
11.8667 | -34.5849 | -19.9407
11.8728 | -34.5854 | -19.9495
11.8789 | -34.5858 | -19.9584
11.885 | -34.5863 | -19.9674
11.8911 | -34.5867 | -19.9762
11.8972 | -34.5872 | -19.9851
11.9033 | -34.5876 | -19.994
11.9094 | -34.5881 | -20.003
11.9155 | -34.5885 | -20.012
11.9216 | -34.589 | -20.0207
11.9277 | -34.5894 | -20.0296
11.9338 | -34.5899 | -20.0387
11.9399 | -34.5903 | -20.0477
11.946 | -34.5908 | -20.0565
11.9521 | -34.5912 | -20.0653
11.9582 | -34.5917 | -20.0742
11.9644 | -34.5921 | -20.083
11.9705 | -34.5926 | -20.0918
11.9766 | -34.593 | -20.1006
11.9827 | -34.5935 | -20.1095
11.9888 | -34.594 | -20.1183
11.9949 | -34.5944 | -20.1268
12.001 | -34.5949 | -20.1358
12.0071 | -34.5953 | -20.1446
12.0132 | -34.5957 | -20.1533
12.0193 | -34.5962 | -20.1624
12.0254 | -34.5966 | -20.1712
12.0315 | -34.5971 | -20.1799
12.0376 | -34.5975 | -20.1887
12.0437 | -34.5979 | -20.1975
12.0498 | -34.5984 | -20.2065
12.0559 | -34.5989 | -20.2152
12.062 | -34.5993 | -20.2238
12.0681 | -34.5998 | -20.2327
12.0742 | -34.6003 | -20.2415
12.0803 | -34.6007 | -20.2505
12.0864 | -34.6011 | -20.2595
12.0925 | -34.6016 | -20.2683
12.0986 | -34.6021 | -20.277
12.1047 | -34.6025 | -20.2859
12.1109 | -34.603 | -20.2946
12.117 | -34.6034 | -20.3034
12.1231 | -34.6039 | -20.3121
12.1292 | -34.6043 | -20.3209
12.1353 | -34.6048 | -20.3296
12.1414 | -34.6052 | -20.3385
12.1475 | -34.6057 | -20.3474
12.1536 | -34.6061 | -20.3563
12.1597 | -34.6066 | -20.3651
12.1658 | -34.6071 | -20.3738
12.1719 | -34.6075 | -20.3824
12.178 | -34.6079 | -20.3912
12.1841 | -34.6084 | -20.4001
12.1902 | -34.6089 | -20.409
12.1963 | -34.6093 | -20.4178
12.2024 | -34.6098 | -20.4264
12.2085 | -34.6102 | -20.4352
12.2146 | -34.6107 | -20.4441
12.2207 | -34.6111 | -20.4529
12.2268 | -34.6116 | -20.4616
12.2329 | -34.6121 | -20.4705
12.239 | -34.6126 | -20.479
12.2451 | -34.613 | -20.4878
12.2513 | -34.6134 | -20.4967
12.2574 | -34.6139 | -20.5056
12.2635 | -34.6143 | -20.5144
12.2696 | -34.6148 | -20.5231
12.2757 | -34.6153 | -20.5316
12.2818 | -34.6157 | -20.5404
12.2879 | -34.6162 | -20.5492
12.294 | -34.6166 | -20.5578
12.3001 | -34.6171 | -20.5666
12.3062 | -34.6176 | -20.5753
12.3123 | -34.618 | -20.5839
12.3184 | -34.6185 | -20.5926
12.3245 | -34.619 | -20.6012
12.3306 | -34.6194 | -20.6099
12.3367 | -34.6198 | -20.6187
12.3428 | -34.6203 | -20.6274
12.3489 | -34.6208 | -20.6358
12.355 | -34.6212 | -20.6447
12.3611 | -34.6217 | -20.6535
12.3672 | -34.6221 | -20.6622
12.3733 | -34.6226 | -20.671
12.3794 | -34.6231 | -20.6795
12.3855 | -34.6235 | -20.6881
12.3916 | -34.624 | -20.6967
12.3978 | -34.6245 | -20.7055
12.4039 | -34.6249 | -20.7142
12.41 | -34.6254 | -20.7229
12.4161 | -34.6259 | -20.7313
12.4222 | -34.6263 | -20.7401
12.4283 | -34.6268 | -20.7487
12.4344 | -34.6272 | -20.7574
12.4405 | -34.6277 | -20.7662
12.4466 | -34.6282 | -20.7749
12.4527 | -34.6286 | -20.7835
12.4588 | -34.6291 | -20.7922
12.4649 | -34.6295 | -20.801
12.471 | -34.63 | -20.8098
12.4771 | -34.6304 | -20.8185
12.4832 | -34.6309 | -20.8271
12.4893 | -34.6313 | -20.8356
12.4954 | -34.6318 | -20.8441
12.5015 | -34.6323 | -20.8528
12.5076 | -34.6327 | -20.8616
12.5137 | -34.6332 | -20.8703
12.5198 | -34.6337 | -20.8789
12.5259 | -34.6341 | -20.8876
12.532 | -34.6346 | -20.8963
12.5382 | -34.635 | -20.9051
12.5443 | -34.6355 | -20.9137
12.5504 | -34.6359 | -20.9224
12.5565 | -34.6364 | -20.931
12.5626 | -34.6369 | -20.9397
12.5687 | -34.6373 | -20.9484
12.5748 | -34.6378 | -20.9571
12.5809 | -34.6382 | -20.9657
12.587 | -34.6387 | -20.9745
12.5931 | -34.6392 | -20.9829
12.5992 | -34.6396 | -20.9916
12.6053 | -34.6401 | -21.0003
12.6114 | -34.6405 | -21.0088
12.6175 | -34.641 | -21.0176
12.6236 | -34.6415 | -21.0262
12.6297 | -34.6419 | -21.0348
12.6358 | -34.6424 | -21.0433
12.6419 | -34.6428 | -21.0519
12.648 | -34.6433 | -21.0606
12.6541 | -34.6438 | -21.0693
12.6602 | -34.6442 | -21.0777
12.6663 | -34.6447 | -21.0863
12.6724 | -34.6452 | -21.0949
12.6785 | -34.6456 | -21.1035
12.6847 | -34.6461 | -21.112
12.6908 | -34.6465 | -21.1207
12.6969 | -34.647 | -21.1292
12.703 | -34.6475 | -21.1379
12.7091 | -34.6479 | -21.1465
12.7152 | -34.6484 | -21.155
12.7213 | -34.6488 | -21.1636
12.7274 | -34.6493 | -21.1723
12.7335 | -34.6498 | -21.1806
12.7396 | -34.6502 | -21.1892
12.7457 | -34.6507 | -21.1979
12.7518 | -34.6511 | -21.2064
12.7579 | -34.6516 | -21.2151
12.764 | -34.6521 | -21.2238
12.7701 | -34.6525 | -21.2324
12.7762 | -34.653 | -21.2409
12.7823 | -34.6535 | -21.2495
12.7884 | -34.654 | -21.2582
12.7945 | -34.6544 | -21.2669
12.8006 | -34.6549 | -21.2754
12.8067 | -34.6554 | -21.2841
12.8128 | -34.6558 | -21.2925
12.8189 | -34.6563 | -21.3011
12.8251 | -34.6567 | -21.3098
12.8312 | -34.6572 | -21.3185
12.8373 | -34.6577 | -21.327
12.8434 | -34.6581 | -21.3357
12.8495 | -34.6586 | -21.3443
12.8556 | -34.659 | -21.3528
12.8617 | -34.6595 | -21.3615
12.8678 | -34.66 | -21.3703
12.8739 | -34.6604 | -21.3787
12.88 | -34.6609 | -21.3872
12.8861 | -34.6613 | -21.3958
12.8922 | -34.6618 | -21.4045
12.8983 | -34.6622 | -21.4131
12.9044 | -34.6627 | -21.4215
12.9105 | -34.6632 | -21.43
12.9166 | -34.6636 | -21.4385
12.9227 | -34.6641 | -21.4471
12.9288 | -34.6645 | -21.4557
12.9349 | -34.665 | -21.4643
12.941 | -34.6655 | -21.4729
12.9471 | -34.666 | -21.4815
12.9532 | -34.6664 | -21.49
12.9593 | -34.6669 | -21.4987
12.9654 | -34.6673 | -21.5074
12.9716 | -34.6678 | -21.5161
12.9777 | -34.6683 | -21.5247
12.9838 | -34.6688 | -21.5332
12.9899 | -34.6692 | -21.5417
12.996 | -34.6697 | -21.5503
13.0021 | -34.6701 | -21.5589
13.0082 | -34.6706 | -21.5674
13.0143 | -34.6711 | -21.5757
13.0204 | -34.6716 | -21.5842
13.0265 | -34.6721 | -21.5927
13.0326 | -34.6725 | -21.6013
13.0387 | -34.673 | -21.6099
13.0448 | -34.6734 | -21.6185
13.0509 | -34.6739 | -21.627
13.057 | -34.6744 | -21.6355
13.0631 | -34.6749 | -21.644
13.0692 | -34.6753 | -21.6525
13.0753 | -34.6758 | -21.6611
13.0814 | -34.6763 | -21.6698
13.0875 | -34.6767 | -21.6782
13.0936 | -34.6772 | -21.6868
13.0997 | -34.6777 | -21.6953
13.1058 | -34.6781 | -21.7037
13.112 | -34.6786 | -21.7123
13.1181 | -34.6791 | -21.7208
13.1242 | -34.6795 | -21.7294
13.1303 | -34.68 | -21.738
13.1364 | -34.6804 | -21.7465
13.1425 | -34.6809 | -21.7551
13.1486 | -34.6813 | -21.7637
13.1547 | -34.6818 | -21.7722
13.1608 | -34.6823 | -21.7807
13.1669 | -34.6827 | -21.7891
13.173 | -34.6832 | -21.7977
13.1791 | -34.6837 | -21.8061
13.1852 | -34.6841 | -21.8146
13.1913 | -34.6846 | -21.8231
13.1974 | -34.685 | -21.8316
13.2035 | -34.6855 | -21.8401
13.2096 | -34.6859 | -21.8485
13.2157 | -34.6864 | -21.8569
13.2218 | -34.6869 | -21.8654
13.2279 | -34.6873 | -21.8738
13.234 | -34.6878 | -21.8822
13.2401 | -34.6883 | -21.8908
13.2462 | -34.6887 | -21.8993
13.2524 | -34.6892 | -21.9079
13.2585 | -34.6897 | -21.9163
13.2646 | -34.6901 | -21.9248
13.2707 | -34.6906 | -21.9333
13.2768 | -34.6911 | -21.942
13.2829 | -34.6915 | -21.9506
13.289 | -34.692 | -21.9589
13.2951 | -34.6925 | -21.9673
13.3012 | -34.6929 | -21.9759
13.3073 | -34.6934 | -21.9846
13.3134 | -34.6939 | -21.9932
13.3195 | -34.6943 | -22.0015
13.3256 | -34.6948 | -22.0099
13.3317 | -34.6953 | -22.0184
13.3378 | -34.6957 | -22.0269
13.3439 | -34.6962 | -22.0355
13.35 | -34.6967 | -22.0441
13.3561 | -34.6971 | -22.0525
13.3622 | -34.6976 | -22.0609
13.3683 | -34.6981 | -22.0693
13.3744 | -34.6986 | -22.0776
13.3805 | -34.699 | -22.0862
13.3866 | -34.6995 | -22.0945
13.3927 | -34.6999 | -22.103
13.3989 | -34.7004 | -22.1113
13.405 | -34.7009 | -22.1198
13.4111 | -34.7013 | -22.1282
13.4172 | -34.7018 | -22.1366
13.4233 | -34.7023 | -22.1452
13.4294 | -34.7027 | -22.1538
13.4355 | -34.7032 | -22.1622
13.4416 | -34.7036 | -22.1707
13.4477 | -34.7041 | -22.1792
13.4538 | -34.7046 | -22.1877
13.4599 | -34.7051 | -22.1961
13.466 | -34.7055 | -22.2047
13.4721 | -34.706 | -22.2131
13.4782 | -34.7064 | -22.2216
13.4843 | -34.7069 | -22.2302
13.4904 | -34.7074 | -22.2386
13.4965 | -34.7078 | -22.2472
13.5026 | -34.7083 | -22.2558
13.5087 | -34.7087 | -22.2642
13.5148 | -34.7092 | -22.2728
13.5209 | -34.7097 | -22.2815
13.527 | -34.7101 | -22.2899
13.5331 | -34.7106 | -22.2983
13.5393 | -34.7111 | -22.3067
13.5454 | -34.7116 | -22.3151
13.5515 | -34.712 | -22.3237
13.5576 | -34.7125 | -22.3321
13.5637 | -34.7129 | -22.3407
13.5698 | -34.7134 | -22.3492
13.5759 | -34.7139 | -22.3574
13.582 | -34.7143 | -22.3658
13.5881 | -34.7148 | -22.3742
13.5942 | -34.7153 | -22.3828
13.6003 | -34.7157 | -22.3912
13.6064 | -34.7162 | -22.3998
13.6125 | -34.7167 | -22.4082
13.6186 | -34.7171 | -22.4166
13.6247 | -34.7176 | -22.425
13.6308 | -34.7181 | -22.4336
13.6369 | -34.7185 | -22.4421
13.643 | -34.719 | -22.4507
13.6491 | -34.7195 | -22.4589
13.6552 | -34.72 | -22.4673
13.6613 | -34.7204 | -22.4756
13.6674 | -34.7209 | -22.4841
13.6735 | -34.7213 | -22.4926
13.6796 | -34.7218 | -22.501
13.6858 | -34.7223 | -22.5093
13.6919 | -34.7228 | -22.5178
13.698 | -34.7232 | -22.5262
13.7041 | -34.7237 | -22.5349
13.7102 | -34.7242 | -22.5433
13.7163 | -34.7247 | -22.5518
13.7224 | -34.7251 | -22.5603
13.7285 | -34.7256 | -22.5688
13.7346 | -34.726 | -22.5772
13.7407 | -34.7265 | -22.5856
13.7468 | -34.727 | -22.5941
13.7529 | -34.7274 | -22.6024
13.759 | -34.7279 | -22.6109
13.7651 | -34.7283 | -22.6194
13.7712 | -34.7288 | -22.6279
13.7773 | -34.7293 | -22.6364
13.7834 | -34.7297 | -22.6447
13.7895 | -34.7302 | -22.6531
13.7956 | -34.7307 | -22.6616
13.8017 | -34.7311 | -22.6699
13.8078 | -34.7316 | -22.6783
13.8139 | -34.732 | -22.6867
13.82 | -34.7325 | -22.6951
13.8262 | -34.733 | -22.7035
13.8323 | -34.7335 | -22.712
13.8384 | -34.7339 | -22.7205
13.8445 | -34.7344 | -22.7289
13.8506 | -34.7349 | -22.7374
13.8567 | -34.7353 | -22.7459
13.8628 | -34.7358 | -22.7544
13.8689 | -34.7363 | -22.7629
13.875 | -34.7367 | -22.7713
13.8811 | -34.7372 | -22.7797
13.8872 | -34.7376 | -22.7881
13.8933 | -34.7381 | -22.7964
13.8994 | -34.7386 | -22.8049
13.9055 | -34.7391 | -22.8134
13.9116 | -34.7395 | -22.8217
13.9177 | -34.74 | -22.8302
13.9238 | -34.7404 | -22.8386
13.9299 | -34.7409 | -22.8469
13.936 | -34.7414 | -22.8554
13.9421 | -34.7418 | -22.8638
13.9482 | -34.7423 | -22.8722
13.9543 | -34.7428 | -22.8806
13.9604 | -34.7432 | -22.8891
13.9665 | -34.7437 | -22.8973
13.9727 | -34.7442 | -22.9058
13.9788 | -34.7447 | -22.9142
13.9849 | -34.7451 | -22.9226
13.991 | -34.7456 | -22.9309
13.9971 | -34.746 | -22.9392
14.0032 | -34.7465 | -22.9477
14.0093 | -34.7469 | -22.9561
14.0154 | -34.7474 | -22.9645
14.0215 | -34.7479 | -22.9728
14.0276 | -34.7484 | -22.9812
14.0337 | -34.7488 | -22.9896
14.0398 | -34.7493 | -22.9979
14.0459 | -34.7498 | -23.0063
14.052 | -34.7502 | -23.0147
14.0581 | -34.7507 | -23.0231
14.0642 | -34.7511 | -23.0316
14.0703 | -34.7516 | -23.04
14.0764 | -34.7521 | -23.0483
14.0825 | -34.7526 | -23.0567
14.0886 | -34.753 | -23.0651
14.0947 | -34.7535 | -23.0736
14.1008 | -34.754 | -23.082
14.1069 | -34.7545 | -23.0902
14.1131 | -34.7549 | -23.0986
14.1192 | -34.7554 | -23.1069
14.1253 | -34.7559 | -23.1152
14.1314 | -34.7563 | -23.1236
14.1375 | -34.7568 | -23.1319
14.1436 | -34.7573 | -23.1402
14.1497 | -34.7577 | -23.1485
14.1558 | -34.7582 | -23.1568
14.1619 | -34.7587 | -23.1651
14.168 | -34.7591 | -23.1735
14.1741 | -34.7596 | -23.1818
14.1802 | -34.7601 | -23.1901
14.1863 | -34.7605 | -23.1984
14.1924 | -34.761 | -23.2068
14.1985 | -34.7615 | -23.2151
14.2046 | -34.762 | -23.2236
14.2107 | -34.7625 | -23.2317
14.2168 | -34.7629 | -23.24
14.2229 | -34.7634 | -23.2484
14.229 | -34.7639 | -23.2568
14.2351 | -34.7643 | -23.2651
14.2412 | -34.7648 | -23.2734
14.2473 | -34.7653 | -23.2817
14.2534 | -34.7657 | -23.2899
14.2596 | -34.7662 | -23.2982
14.2657 | -34.7666 | -23.3065
14.2718 | -34.7671 | -23.3149
14.2779 | -34.7676 | -23.3233
14.284 | -34.7681 | -23.3316
14.2901 | -34.7685 | -23.34
14.2962 | -34.769 | -23.3484
14.3023 | -34.7695 | -23.3568
14.3084 | -34.77 | -23.3651
14.3145 | -34.7704 | -23.3734
14.3206 | -34.7709 | -23.3819
14.3267 | -34.7714 | -23.3903
14.3328 | -34.7719 | -23.3985
14.3389 | -34.7723 | -23.4068
14.345 | -34.7728 | -23.4152
14.3511 | -34.7732 | -23.4234
14.3572 | -34.7737 | -23.4318
14.3633 | -34.7742 | -23.4402
14.3694 | -34.7747 | -23.4484
14.3755 | -34.7752 | -23.4568
14.3816 | -34.7756 | -23.4653
14.3877 | -34.7761 | -23.4734
14.3938 | -34.7766 | -23.4819
14.4 | -34.7771 | -23.49
14.4061 | -34.7775 | -23.4983
14.4122 | -34.778 | -23.5066
14.4183 | -34.7785 | -23.5148
14.4244 | -34.7789 | -23.5231
14.4305 | -34.7794 | -23.5315
14.4366 | -34.7799 | -23.5397
14.4427 | -34.7804 | -23.5482
14.4488 | -34.7809 | -23.5566
14.4549 | -34.7814 | -23.5647
14.461 | -34.7818 | -23.5728
14.4671 | -34.7823 | -23.5812
14.4732 | -34.7828 | -23.5896
14.4793 | -34.7832 | -23.598
14.4854 | -34.7837 | -23.6063
14.4915 | -34.7842 | -23.6143
14.4976 | -34.7847 | -23.6225
14.5037 | -34.7852 | -23.6308
14.5098 | -34.7857 | -23.639
14.5159 | -34.7861 | -23.6474
14.522 | -34.7866 | -23.6556
14.5281 | -34.7871 | -23.6637
14.5342 | -34.7876 | -23.6721
14.5403 | -34.788 | -23.6803
14.5465 | -34.7885 | -23.6886
14.5526 | -34.789 | -23.697
14.5587 | -34.7894 | -23.7053
14.5648 | -34.7899 | -23.7134
14.5709 | -34.7904 | -23.7217
14.577 | -34.7909 | -23.7299
14.5831 | -34.7913 | -23.7382
14.5892 | -34.7918 | -23.7465
14.5953 | -34.7923 | -23.7545
14.6014 | -34.7928 | -23.7629
14.6075 | -34.7932 | -23.7711
14.6136 | -34.7937 | -23.7794
14.6197 | -34.7942 | -23.7878
14.6258 | -34.7946 | -23.796
14.6319 | -34.7951 | -23.8041
14.638 | -34.7956 | -23.8124
14.6441 | -34.796 | -23.8208
14.6502 | -34.7965 | -23.8291
14.6563 | -34.797 | -23.8373
14.6624 | -34.7975 | -23.8456
14.6685 | -34.798 | -23.8536
14.6746 | -34.7984 | -23.8619
14.6807 | -34.7989 | -23.87
14.6869 | -34.7994 | -23.8782
14.693 | -34.7999 | -23.8864
14.6991 | -34.8003 | -23.8946
14.7052 | -34.8008 | -23.9027
14.7113 | -34.8013 | -23.9111
14.7174 | -34.8018 | -23.9193
14.7235 | -34.8022 | -23.9274
14.7296 | -34.8027 | -23.9356
14.7357 | -34.8032 | -23.9438
14.7418 | -34.8037 | -23.952
14.7479 | -34.8042 | -23.9602
14.754 | -34.8046 | -23.9685
14.7601 | -34.8051 | -23.9767
14.7662 | -34.8056 | -23.9848
14.7723 | -34.8061 | -23.9928
14.7784 | -34.8065 | -24.0011
14.7845 | -34.807 | -24.0094
14.7906 | -34.8075 | -24.0177
14.7967 | -34.808 | -24.0258
14.8028 | -34.8085 | -24.0341
14.8089 | -34.809 | -24.0422
14.815 | -34.8094 | -24.0504
14.8211 | -34.8099 | -24.0586
14.8272 | -34.8104 | -24.0669
14.8334 | -34.8109 | -24.075
14.8395 | -34.8114 | -24.0832
14.8456 | -34.8118 | -24.0911
14.8517 | -34.8123 | -24.0993
14.8578 | -34.8128 | -24.1075
14.8639 | -34.8133 | -24.1158
14.87 | -34.8138 | -24.124
14.8761 | -34.8143 | -24.1322
14.8822 | -34.8148 | -24.1405
14.8883 | -34.8153 | -24.1487
14.8944 | -34.8157 | -24.1568
14.9005 | -34.8162 | -24.165
14.9066 | -34.8167 | -24.1732
14.9127 | -34.8172 | -24.1813
14.9188 | -34.8176 | -24.1896
14.9249 | -34.8181 | -24.1977
14.931 | -34.8186 | -24.2058
14.9371 | -34.8191 | -24.2139
14.9432 | -34.8195 | -24.2221
14.9493 | -34.82 | -24.2302
14.9554 | -34.8205 | -24.2384
14.9615 | -34.821 | -24.2467
14.9676 | -34.8215 | -24.255
14.9738 | -34.8219 | -24.2631
14.9799 | -34.8224 | -24.2713
14.986 | -34.8229 | -24.2793
14.9921 | -34.8234 | -24.2875
14.9982 | -34.8239 | -24.2957
15.0043 | -34.8243 | -24.3038
15.0104 | -34.8248 | -24.312
15.0165 | -34.8253 | -24.3202
15.0226 | -34.8258 | -24.3281
15.0287 | -34.8263 | -24.3363
15.0348 | -34.8268 | -24.3445
15.0409 | -34.8273 | -24.3527
15.047 | -34.8278 | -24.3609
15.0531 | -34.8283 | -24.3689
15.0592 | -34.8288 | -24.3772
15.0653 | -34.8292 | -24.3854
15.0714 | -34.8297 | -24.3935
15.0775 | -34.8302 | -24.4016
15.0836 | -34.8307 | -24.4097
15.0897 | -34.8312 | -24.4175
15.0958 | -34.8317 | -24.4257
15.1019 | -34.8322 | -24.434
15.108 | -34.8326 | -24.442
15.1141 | -34.8331 | -24.4502
15.1203 | -34.8336 | -24.4584
15.1264 | -34.8341 | -24.4665
15.1325 | -34.8346 | -24.4747
15.1386 | -34.8351 | -24.4828
15.1447 | -34.8356 | -24.4909
15.1508 | -34.8361 | -24.4991
15.1569 | -34.8366 | -24.5072
15.163 | -34.8371 | -24.5151
15.1691 | -34.8376 | -24.5231
15.1752 | -34.838 | -24.5311
15.1813 | -34.8385 | -24.5392
15.1874 | -34.839 | -24.5474
15.1935 | -34.8395 | -24.5553
15.1996 | -34.84 | -24.5632
15.2057 | -34.8405 | -24.5712
15.2118 | -34.841 | -24.5793
15.2179 | -34.8415 | -24.5875
15.224 | -34.842 | -24.5956
15.2301 | -34.8425 | -24.6034
15.2362 | -34.843 | -24.6115
15.2423 | -34.8434 | -24.6196
15.2484 | -34.8439 | -24.6277
15.2545 | -34.8444 | -24.636
15.2607 | -34.8449 | -24.644
15.2668 | -34.8454 | -24.6519
15.2729 | -34.8459 | -24.6599
15.279 | -34.8464 | -24.668
15.2851 | -34.8469 | -24.676
15.2912 | -34.8474 | -24.6841
15.2973 | -34.8479 | -24.6922
15.3034 | -34.8484 | -24.7002
15.3095 | -34.8489 | -24.7083
15.3156 | -34.8494 | -24.7163
15.3217 | -34.8499 | -24.7244
15.3278 | -34.8504 | -24.7325
15.3339 | -34.8509 | -24.7403
15.34 | -34.8514 | -24.7483
15.3461 | -34.8519 | -24.7562
15.3522 | -34.8523 | -24.7643
15.3583 | -34.8528 | -24.7726
15.3644 | -34.8533 | -24.7805
15.3705 | -34.8538 | -24.7883
15.3766 | -34.8543 | -24.7964
15.3827 | -34.8548 | -24.8043
15.3888 | -34.8553 | -24.8123
15.3949 | -34.8558 | -24.8204
15.401 | -34.8563 | -24.8286
15.4072 | -34.8568 | -24.8364
15.4133 | -34.8572 | -24.8445
15.4194 | -34.8577 | -24.8526
15.4255 | -34.8582 | -24.8607
15.4316 | -34.8587 | -24.8687
15.4377 | -34.8592 | -24.8769
15.4438 | -34.8597 | -24.8848
15.4499 | -34.8602 | -24.8929
15.456 | -34.8607 | -24.9007
15.4621 | -34.8612 | -24.9087
15.4682 | -34.8617 | -24.9169
15.4743 | -34.8622 | -24.9248
15.4804 | -34.8627 | -24.9328
15.4865 | -34.8632 | -24.9408
15.4926 | -34.8637 | -24.949
15.4987 | -34.8642 | -24.9571
15.5048 | -34.8647 | -24.9651
15.5109 | -34.8652 | -24.973
15.517 | -34.8657 | -24.981
15.5231 | -34.8662 | -24.989
15.5292 | -34.8667 | -24.9969
15.5353 | -34.8671 | -25.005
15.5414 | -34.8676 | -25.0129
15.5476 | -34.8682 | -25.0207
15.5537 | -34.8687 | -25.0287
15.5598 | -34.8692 | -25.0368
15.5659 | -34.8696 | -25.0448
15.572 | -34.8702 | -25.0528
15.5781 | -34.8706 | -25.0607
15.5842 | -34.8712 | -25.0685
15.5903 | -34.8717 | -25.0765
15.5964 | -34.8722 | -25.0845
15.6025 | -34.8727 | -25.0924
15.6086 | -34.8732 | -25.1005
15.6147 | -34.8736 | -25.1086
15.6208 | -34.8742 | -25.1165
15.6269 | -34.8747 | -25.1244
15.633 | -34.8752 | -25.1324
15.6391 | -34.8757 | -25.1405
15.6452 | -34.8761 | -25.1485
15.6513 | -34.8767 | -25.1563
15.6574 | -34.8772 | -25.1644
15.6635 | -34.8777 | -25.1723
15.6696 | -34.8782 | -25.1804
15.6757 | -34.8787 | -25.1884
15.6818 | -34.8792 | -25.1964
15.688 | -34.8797 | -25.2043
15.6941 | -34.8802 | -25.2121
15.7002 | -34.8806 | -25.2201
15.7063 | -34.8811 | -25.228
15.7124 | -34.8816 | -25.2361
15.7185 | -34.8821 | -25.244
15.7246 | -34.8826 | -25.2519
15.7307 | -34.8831 | -25.2598
15.7368 | -34.8836 | -25.2678
15.7429 | -34.8842 | -25.2757
15.749 | -34.8847 | -25.2837
15.7551 | -34.8852 | -25.2917
15.7612 | -34.8857 | -25.2995
15.7673 | -34.8862 | -25.3074
15.7734 | -34.8867 | -25.3154
15.7795 | -34.8872 | -25.3233
15.7856 | -34.8877 | -25.3313
15.7917 | -34.8882 | -25.3391
15.7978 | -34.8887 | -25.3471
15.8039 | -34.8892 | -25.3551
15.81 | -34.8897 | -25.363
15.8161 | -34.8901 | -25.3709
15.8222 | -34.8906 | -25.3787
15.8283 | -34.8912 | -25.3865
15.8345 | -34.8917 | -25.3944
15.8406 | -34.8922 | -25.4023
15.8467 | -34.8927 | -25.4102
15.8528 | -34.8932 | -25.4182
15.8589 | -34.8937 | -25.4262
15.865 | -34.8943 | -25.4339
15.8711 | -34.8948 | -25.4419
15.8772 | -34.8953 | -25.4497
15.8833 | -34.8958 | -25.4577
15.8894 | -34.8963 | -25.4658
15.8955 | -34.8968 | -25.4737
15.9016 | -34.8973 | -25.4813
15.9077 | -34.8977 | -25.4893
15.9138 | -34.8982 | -25.4972
15.9199 | -34.8987 | -25.505
15.926 | -34.8992 | -25.5128
15.9321 | -34.8997 | -25.5206
15.9382 | -34.9003 | -25.5284
15.9443 | -34.9008 | -25.5363
15.9504 | -34.9013 | -25.5443
15.9565 | -34.9018 | -25.5521
15.9626 | -34.9022 | -25.5601
15.9687 | -34.9028 | -25.5678
15.9749 | -34.9032 | -25.5756
15.981 | -34.9037 | -25.5835
15.9871 | -34.9042 | -25.5915
15.9932 | -34.9047 | -25.5993
15.9993 | -34.9052 | -25.6072
16.0054 | -34.9058 | -25.6148
16.0115 | -34.9063 | -25.6228
16.0176 | -34.9068 | -25.6306
16.0237 | -34.9073 | -25.6386
16.0298 | -34.9078 | -25.6463
16.0359 | -34.9083 | -25.6542
16.042 | -34.9088 | -25.6619
16.0481 | -34.9093 | -25.6699
16.0542 | -34.9098 | -25.6777
16.0603 | -34.9103 | -25.6856
16.0664 | -34.9108 | -25.6934
16.0725 | -34.9113 | -25.7011
16.0786 | -34.9118 | -25.7089
16.0847 | -34.9123 | -25.7169
16.0908 | -34.9128 | -25.7248
16.0969 | -34.9133 | -25.7328
16.103 | -34.9138 | -25.7406
16.1091 | -34.9144 | -25.7482
16.1152 | -34.9149 | -25.756
16.1214 | -34.9154 | -25.764
16.1275 | -34.9159 | -25.7718
16.1336 | -34.9164 | -25.7796
16.1397 | -34.9169 | -25.7873
16.1458 | -34.9174 | -25.7951
16.1519 | -34.918 | -25.8029
16.158 | -34.9184 | -25.8108
16.1641 | -34.9189 | -25.8188
16.1702 | -34.9195 | -25.8265
16.1763 | -34.92 | -25.8344
16.1824 | -34.9205 | -25.8422
16.1885 | -34.921 | -25.8499
16.1946 | -34.9215 | -25.8576
16.2007 | -34.922 | -25.8654
16.2068 | -34.9225 | -25.8732
16.2129 | -34.923 | -25.8811
16.219 | -34.9236 | -25.8886
16.2251 | -34.9241 | -25.8962
16.2312 | -34.9246 | -25.904
16.2373 | -34.9251 | -25.9117
16.2434 | -34.9256 | -25.9195
16.2495 | -34.9261 | -25.9271
16.2556 | -34.9266 | -25.935
16.2618 | -34.9271 | -25.9427
16.2679 | -34.9276 | -25.9506
16.274 | -34.9281 | -25.9585
16.2801 | -34.9286 | -25.9662
16.2862 | -34.9292 | -25.974
16.2923 | -34.9297 | -25.9818
16.2984 | -34.9302 | -25.9897
16.3045 | -34.9307 | -25.9975
16.3106 | -34.9312 | -26.0053
16.3167 | -34.9317 | -26.0131
16.3228 | -34.9322 | -26.0207
16.3289 | -34.9327 | -26.0285
16.335 | -34.9332 | -26.0362
16.3411 | -34.9337 | -26.044
16.3472 | -34.9342 | -26.0519
16.3533 | -34.9347 | -26.0596
16.3594 | -34.9352 | -26.0671
16.3655 | -34.9357 | -26.075
16.3716 | -34.9362 | -26.0827
16.3777 | -34.9367 | -26.0904
16.3838 | -34.9372 | -26.0981
16.3899 | -34.9377 | -26.1059
16.396 | -34.9382 | -26.1137
16.4021 | -34.9387 | -26.1213
16.4083 | -34.9392 | -26.129
16.4144 | -34.9397 | -26.1368
16.4205 | -34.9402 | -26.1447
16.4266 | -34.9407 | -26.1523
16.4327 | -34.9413 | -26.1601
16.4388 | -34.9417 | -26.1679
16.4449 | -34.9423 | -26.1756
16.451 | -34.9428 | -26.1835
16.4571 | -34.9433 | -26.1914
16.4632 | -34.9438 | -26.1989
16.4693 | -34.9443 | -26.2068
16.4754 | -34.9448 | -26.2146
16.4815 | -34.9453 | -26.2224
16.4876 | -34.9458 | -26.2301
16.4937 | -34.9463 | -26.2378
16.4998 | -34.9468 | -26.2453
16.5059 | -34.9473 | -26.253
16.512 | -34.9478 | -26.2609
16.5181 | -34.9483 | -26.2687
16.5242 | -34.9488 | -26.2766
16.5303 | -34.9493 | -26.2842
16.5364 | -34.9498 | -26.2919
16.5425 | -34.9503 | -26.2996
16.5487 | -34.9508 | -26.3074
16.5548 | -34.9513 | -26.3152
16.5609 | -34.9518 | -26.3229
16.567 | -34.9523 | -26.3305
16.5731 | -34.9529 | -26.3383
16.5792 | -34.9534 | -26.3461
16.5853 | -34.9539 | -26.3539
16.5914 | -34.9544 | -26.3617
16.5975 | -34.9549 | -26.3695
16.6036 | -34.9554 | -26.3771
16.6097 | -34.9559 | -26.3849
16.6158 | -34.9564 | -26.3927
16.6219 | -34.9569 | -26.4004
16.628 | -34.9574 | -26.4081
16.6341 | -34.9579 | -26.4159
16.6402 | -34.9584 | -26.4234
16.6463 | -34.9589 | -26.4312
16.6524 | -34.9594 | -26.4387
16.6585 | -34.9599 | -26.4465
16.6646 | -34.9605 | -26.4542
16.6707 | -34.961 | -26.4618
16.6768 | -34.9615 | -26.4695
16.6829 | -34.962 | -26.4771
16.689 | -34.9625 | -26.4848
16.6952 | -34.963 | -26.4926
16.7013 | -34.9635 | -26.5003
16.7074 | -34.964 | -26.5079
16.7135 | -34.9645 | -26.5157
16.7196 | -34.965 | -26.5234
16.7257 | -34.9655 | -26.5311
16.7318 | -34.966 | -26.5387
16.7379 | -34.9665 | -26.5464
16.744 | -34.967 | -26.554
16.7501 | -34.9675 | -26.5617
16.7562 | -34.968 | -26.5694
16.7623 | -34.9685 | -26.5769
16.7684 | -34.969 | -26.5846
16.7745 | -34.9695 | -26.5924
16.7806 | -34.97 | -26.5999
16.7867 | -34.9705 | -26.6077
16.7928 | -34.971 | -26.6154
16.7989 | -34.9715 | -26.623
16.805 | -34.972 | -26.6307
16.8111 | -34.9725 | -26.6383
16.8172 | -34.973 | -26.6461
16.8233 | -34.9736 | -26.6539
16.8294 | -34.9741 | -26.6615
16.8356 | -34.9746 | -26.6692
16.8417 | -34.9751 | -26.6768
16.8478 | -34.9756 | -26.6844
16.8539 | -34.9761 | -26.6919
16.86 | -34.9766 | -26.6997
16.8661 | -34.9771 | -26.7073
16.8722 | -34.9776 | -26.715
16.8783 | -34.9781 | -26.7227
16.8844 | -34.9786 | -26.7301
16.8905 | -34.9791 | -26.7378
16.8966 | -34.9796 | -26.7455
16.9027 | -34.9801 | -26.7531
16.9088 | -34.9806 | -26.7608
16.9149 | -34.9811 | -26.7685
16.921 | -34.9816 | -26.7759
16.9271 | -34.9822 | -26.7837
16.9332 | -34.9826 | -26.7913
16.9393 | -34.9831 | -26.799
16.9454 | -34.9836 | -26.8066
16.9515 | -34.9841 | -26.8143
16.9576 | -34.9846 | -26.8217
16.9637 | -34.9852 | -26.8293
16.9698 | -34.9857 | -26.8368
16.9759 | -34.9861 | -26.8444
16.9821 | -34.9867 | -26.8521
16.9882 | -34.9872 | -26.8596
16.9943 | -34.9877 | -26.8672
17.0004 | -34.9882 | -26.8749
17.0065 | -34.9887 | -26.8825
17.0126 | -34.9892 | -26.8902
17.0187 | -34.9897 | -26.898
17.0248 | -34.9902 | -26.9054
17.0309 | -34.9907 | -26.913
17.037 | -34.9912 | -26.9206
17.0431 | -34.9917 | -26.9282
17.0492 | -34.9922 | -26.9359
17.0553 | -34.9927 | -26.9435
17.0614 | -34.9932 | -26.951
17.0675 | -34.9937 | -26.9586
17.0736 | -34.9942 | -26.9663
17.0797 | -34.9947 | -26.9737
17.0858 | -34.9952 | -26.9815
17.0919 | -34.9957 | -26.9891
17.098 | -34.9962 | -26.9966
17.1041 | -34.9968 | -27.0041
17.1102 | -34.9973 | -27.0117
17.1163 | -34.9978 | -27.0193
17.1225 | -34.9983 | -27.0268
17.1286 | -34.9988 | -27.0343
17.1347 | -34.9993 | -27.0419
17.1408 | -34.9998 | -27.0495
17.1469 | -35.0003 | -27.0571
17.153 | -35.0008 | -27.0647
17.1591 | -35.0013 | -27.0722
17.1652 | -35.0018 | -27.0795
17.1713 | -35.0023 | -27.087
17.1774 | -35.0028 | -27.0946
17.1835 | -35.0033 | -27.1021
17.1896 | -35.0037 | -27.1096
17.1957 | -35.0043 | -27.1172
17.2018 | -35.0048 | -27.1245
17.2079 | -35.0053 | -27.132
17.214 | -35.0058 | -27.1397
17.2201 | -35.0063 | -27.1473
17.2262 | -35.0068 | -27.155
17.2323 | -35.0073 | -27.1626
17.2384 | -35.0078 | -27.1701
17.2445 | -35.0083 | -27.1776
17.2506 | -35.0088 | -27.1851
17.2567 | -35.0093 | -27.1928
17.2628 | -35.0098 | -27.2003
17.269 | -35.0103 | -27.2078
17.2751 | -35.0108 | -27.2153
17.2812 | -35.0113 | -27.223
17.2873 | -35.0118 | -27.2306
17.2934 | -35.0123 | -27.2381
17.2995 | -35.0128 | -27.2458
17.3056 | -35.0133 | -27.2532
17.3117 | -35.0138 | -27.2609
17.3178 | -35.0143 | -27.2684
17.3239 | -35.0148 | -27.2759
17.33 | -35.0153 | -27.2834
17.3361 | -35.0158 | -27.291
17.3422 | -35.0163 | -27.2983
17.3483 | -35.0167 | -27.3059
17.3544 | -35.0173 | -27.3135
17.3605 | -35.0177 | -27.3209
17.3666 | -35.0182 | -27.3284
17.3727 | -35.0187 | -27.336
17.3788 | -35.0192 | -27.3434
17.3849 | -35.0197 | -27.3508
17.391 | -35.0202 | -27.3584
17.3971 | -35.0207 | -27.3659
17.4032 | -35.0212 | -27.3734
17.4094 | -35.0217 | -27.3808
17.4155 | -35.0222 | -27.3882
17.4216 | -35.0227 | -27.3957
17.4277 | -35.0232 | -27.4032
17.4338 | -35.0237 | -27.4107
17.4399 | -35.0242 | -27.4183
17.446 | -35.0247 | -27.4256
17.4521 | -35.0252 | -27.4332
17.4582 | -35.0257 | -27.4406
17.4643 | -35.0262 | -27.448
17.4704 | -35.0267 | -27.4556
17.4765 | -35.0272 | -27.4632
17.4826 | -35.0277 | -27.4706
17.4887 | -35.0282 | -27.4781
17.4948 | -35.0287 | -27.4857
17.5009 | -35.0292 | -27.4932
17.507 | -35.0297 | -27.5008
17.5131 | -35.0302 | -27.5082
17.5192 | -35.0307 | -27.5155
17.5253 | -35.0312 | -27.523
17.5314 | -35.0317 | -27.5305
17.5375 | -35.0322 | -27.538
17.5436 | -35.0326 | -27.5456
17.5497 | -35.0331 | -27.5532
17.5559 | -35.0336 | -27.5606
17.562 | -35.0341 | -27.5681
17.5681 | -35.0346 | -27.5755
17.5742 | -35.0351 | -27.5831
17.5803 | -35.0356 | -27.5905
17.5864 | -35.0361 | -27.5979
17.5925 | -35.0366 | -27.6054
17.5986 | -35.0371 | -27.6128
17.6047 | -35.0376 | -27.6205
17.6108 | -35.0381 | -27.6279
17.6169 | -35.0386 | -27.6354
17.623 | -35.0391 | -27.6427
17.6291 | -35.0396 | -27.6502
17.6352 | -35.0401 | -27.6577
17.6413 | -35.0406 | -27.6653
17.6474 | -35.0411 | -27.6727
17.6535 | -35.0416 | -27.68
17.6596 | -35.0421 | -27.6874
17.6657 | -35.0426 | -27.6949
17.6718 | -35.0431 | -27.7023
17.6779 | -35.0435 | -27.7097
17.684 | -35.0441 | -27.7172
17.6901 | -35.0445 | -27.7247
17.6963 | -35.045 | -27.732
17.7024 | -35.0455 | -27.7394
17.7085 | -35.046 | -27.7468
17.7146 | -35.0465 | -27.7542
17.7207 | -35.047 | -27.7615
17.7268 | -35.0475 | -27.7688
17.7329 | -35.048 | -27.7763
17.739 | -35.0485 | -27.7837
17.7451 | -35.049 | -27.7911
17.7512 | -35.0495 | -27.7986
17.7573 | -35.05 | -27.8061
17.7634 | -35.0505 | -27.8133
17.7695 | -35.051 | -27.8208
17.7756 | -35.0515 | -27.8281
17.7817 | -35.052 | -27.8356
17.7878 | -35.0525 | -27.843
17.7939 | -35.053 | -27.8505
17.8 | -35.0535 | -27.8578
17.8061 | -35.054 | -27.8653
17.8122 | -35.0545 | -27.8727
17.8183 | -35.055 | -27.8801
17.8244 | -35.0555 | -27.8875
17.8305 | -35.056 | -27.895
17.8366 | -35.0565 | -27.9022
17.8428 | -35.0569 | -27.9096
17.8489 | -35.0574 | -27.9169
17.855 | -35.0579 | -27.9242
17.8611 | -35.0584 | -27.9315
17.8672 | -35.0589 | -27.9387
17.8733 | -35.0594 | -27.9463
17.8794 | -35.0599 | -27.9537
17.8855 | -35.0604 | -27.9612
17.8916 | -35.0609 | -27.9686
17.8977 | -35.0614 | -27.9761
17.9038 | -35.0619 | -27.9832
17.9099 | -35.0624 | -27.9905
17.916 | -35.0629 | -27.998
17.9221 | -35.0634 | -28.0054
17.9282 | -35.0639 | -28.0129
17.9343 | -35.0644 | -28.0203
17.9404 | -35.0649 | -28.0275
17.9465 | -35.0654 | -28.0349
17.9526 | -35.0658 | -28.0423
17.9587 | -35.0663 | -28.0498
17.9648 | -35.0668 | -28.0571
17.9709 | -35.0673 | -28.0645
17.977 | -35.0678 | -28.0718
17.9832 | -35.0683 | -28.0793
17.9893 | -35.0688 | -28.0867
17.9954 | -35.0693 | -28.0942
18.0015 | -35.0697 | -28.1016
18.0076 | -35.0703 | -28.1088
18.0137 | -35.0707 | -28.1162
18.0198 | -35.0712 | -28.1235
18.0259 | -35.0717 | -28.1309
18.032 | -35.0722 | -28.1382
18.0381 | -35.0727 | -28.1456
18.0442 | -35.0732 | -28.1529
18.0503 | -35.0737 | -28.1603
18.0564 | -35.0741 | -28.1676
18.0625 | -35.0746 | -28.1751
18.0686 | -35.0751 | -28.1824
18.0747 | -35.0756 | -28.1898
18.0808 | -35.0761 | -28.1971
18.0869 | -35.0766 | -28.2045
18.093 | -35.0771 | -28.2119
18.0991 | -35.0776 | -28.2193
18.1052 | -35.078 | -28.2267
18.1113 | -35.0785 | -28.234
18.1174 | -35.079 | -28.2412
18.1236 | -35.0795 | -28.2486
18.1297 | -35.08 | -28.256
18.1358 | -35.0805 | -28.2634
18.1419 | -35.081 | -28.2708
18.148 | -35.0815 | -28.278
18.1541 | -35.082 | -28.2851
18.1602 | -35.0825 | -28.2924
18.1663 | -35.0829 | -28.2997
18.1724 | -35.0834 | -28.3071
18.1785 | -35.0839 | -28.3144
18.1846 | -35.0844 | -28.3216
18.1907 | -35.0849 | -28.329
18.1968 | -35.0853 | -28.3364
18.2029 | -35.0858 | -28.3438
18.209 | -35.0863 | -28.3512
18.2151 | -35.0868 | -28.3586
18.2212 | -35.0873 | -28.3658
18.2273 | -35.0878 | -28.3731
18.2334 | -35.0883 | -28.3805
18.2395 | -35.0888 | -28.3879
18.2456 | -35.0893 | -28.3952
18.2517 | -35.0897 | -28.4026
18.2578 | -35.0902 | -28.4098
18.2639 | -35.0907 | -28.4172
18.2701 | -35.0912 | -28.4246
18.2762 | -35.0917 | -28.4318
18.2823 | -35.0922 | -28.4391
18.2884 | -35.0927 | -28.4462
18.2945 | -35.0932 | -28.4533
18.3006 | -35.0937 | -28.4607
18.3067 | -35.0942 | -28.468
18.3128 | -35.0947 | -28.4754
18.3189 | -35.0951 | -28.4827
18.325 | -35.0956 | -28.4898
18.3311 | -35.0961 | -28.4971
18.3372 | -35.0966 | -28.5045
18.3433 | -35.0971 | -28.5119
18.3494 | -35.0975 | -28.5192
18.3555 | -35.0981 | -28.5266
18.3616 | -35.0986 | -28.5336
18.3677 | -35.099 | -28.541
18.3738 | -35.0995 | -28.5482
18.3799 | -35.1 | -28.5555
18.386 | -35.1005 | -28.5627
18.3921 | -35.101 | -28.57
18.3982 | -35.1015 | -28.577
18.4043 | -35.1019 | -28.5843
18.4105 | -35.1024 | -28.5917
18.4166 | -35.1029 | -28.5989
18.4227 | -35.1034 | -28.6062
18.4288 | -35.1039 | -28.6136
18.4349 | -35.1044 | -28.6207
18.441 | -35.1048 | -28.6278
18.4471 | -35.1053 | -28.6351
18.4532 | -35.1058 | -28.6424
18.4593 | -35.1063 | -28.6497
18.4654 | -35.1068 | -28.6569
18.4715 | -35.1073 | -28.6642
18.4776 | -35.1077 | -28.6716
18.4837 | -35.1082 | -28.6789
18.4898 | -35.1087 | -28.6862
18.4959 | -35.1092 | -28.6935
18.502 | -35.1097 | -28.7006
18.5081 | -35.1101 | -28.7079
18.5142 | -35.1106 | -28.7152
18.5203 | -35.1111 | -28.7225
18.5264 | -35.1116 | -28.7297
18.5325 | -35.1121 | -28.737
18.5386 | -35.1126 | -28.7441
18.5447 | -35.1131 | -28.7514
18.5508 | -35.1136 | -28.7586
18.557 | -35.114 | -28.7659
18.5631 | -35.1145 | -28.7731
18.5692 | -35.115 | -28.7804
18.5753 | -35.1155 | -28.7875
18.5814 | -35.116 | -28.7948
18.5875 | -35.1164 | -28.8021
18.5936 | -35.1169 | -28.8093
18.5997 | -35.1174 | -28.8166
18.6058 | -35.1178 | -28.8237
18.6119 | -35.1183 | -28.8309
18.618 | -35.1188 | -28.8381
18.6241 | -35.1193 | -28.8454
18.6302 | -35.1198 | -28.8528
18.6363 | -35.1203 | -28.86
18.6424 | -35.1207 | -28.8671
18.6485 | -35.1212 | -28.8743
18.6546 | -35.1217 | -28.8814
18.6607 | -35.1222 | -28.8886
18.6668 | -35.1227 | -28.8958
18.6729 | -35.1232 | -28.903
18.679 | -35.1236 | -28.9099
18.6851 | -35.1241 | -28.9171
18.6912 | -35.1246 | -28.9242
18.6974 | -35.1251 | -28.9314
18.7035 | -35.1256 | -28.9386
18.7096 | -35.1261 | -28.946
18.7157 | -35.1265 | -28.9529
18.7218 | -35.127 | -28.9602
18.7279 | -35.1275 | -28.9674
18.734 | -35.128 | -28.9748
18.7401 | -35.1285 | -28.9819
18.7462 | -35.129 | -28.9889
18.7523 | -35.1295 | -28.9961
18.7584 | -35.13 | -29.0034
18.7645 | -35.1305 | -29.0105
18.7706 | -35.1309 | -29.0178
18.7767 | -35.1314 | -29.0249
18.7828 | -35.1319 | -29.0319
18.7889 | -35.1324 | -29.0391
18.795 | -35.1329 | -29.0464
18.8011 | -35.1333 | -29.0535
18.8072 | -35.1338 | -29.0609
18.8133 | -35.1343 | -29.068
18.8194 | -35.1348 | -29.0749
18.8255 | -35.1353 | -29.0821
18.8316 | -35.1358 | -29.0893
18.8377 | -35.1362 | -29.0965
18.8439 | -35.1367 | -29.1037
18.85 | -35.1372 | -29.111
18.8561 | -35.1377 | -29.1179
18.8622 | -35.1382 | -29.1252
18.8683 | -35.1386 | -29.1324
18.8744 | -35.1391 | -29.1395
18.8805 | -35.1396 | -29.1468
18.8866 | -35.1401 | -29.1541
18.8927 | -35.1406 | -29.1611
18.8988 | -35.141 | -29.1682
18.9049 | -35.1415 | -29.1752
18.911 | -35.142 | -29.1823
18.9171 | -35.1425 | -29.1894
18.9232 | -35.143 | -29.1964
18.9293 | -35.1435 | -29.2036
18.9354 | -35.1439 | -29.2108
18.9415 | -35.1444 | -29.218
18.9476 | -35.1449 | -29.2251
18.9537 | -35.1454 | -29.2323
18.9598 | -35.1458 | -29.2392
18.9659 | -35.1463 | -29.2463
18.972 | -35.1468 | -29.2535
18.9781 | -35.1473 | -29.2607
18.9843 | -35.1477 | -29.2679
18.9904 | -35.1482 | -29.275
18.9965 | -35.1487 | -29.2819
19.0026 | -35.1492 | -29.2892
19.0087 | -35.1497 | -29.2964
19.0148 | -35.1501 | -29.3035
19.0209 | -35.1506 | -29.3108
19.027 | -35.1511 | -29.3181
19.0331 | -35.1516 | -29.3251
19.0392 | -35.1521 | -29.3323
19.0453 | -35.1526 | -29.3394
19.0514 | -35.1531 | -29.3465
19.0575 | -35.1535 | -29.3536
19.0636 | -35.154 | -29.3605
19.0697 | -35.1545 | -29.3677
19.0758 | -35.155 | -29.3749
19.0819 | -35.1555 | -29.382
19.088 | -35.1559 | -29.389
19.0941 | -35.1564 | -29.3961
19.1002 | -35.1569 | -29.4031
19.1063 | -35.1574 | -29.4101
19.1124 | -35.1578 | -29.4172
19.1185 | -35.1583 | -29.4243
19.1246 | -35.1588 | -29.4313
19.1308 | -35.1593 | -29.4385
19.1369 | -35.1597 | -29.4453
19.143 | -35.1602 | -29.4525
19.1491 | -35.1607 | -29.4596
19.1552 | -35.1612 | -29.4667
19.1613 | -35.1617 | -29.4739
19.1674 | -35.1621 | -29.481
19.1735 | -35.1626 | -29.488
19.1796 | -35.1631 | -29.495
19.1857 | -35.1636 | -29.5021
19.1918 | -35.1641 | -29.5092
19.1979 | -35.1646 | -29.5163
19.204 | -35.165 | -29.5233
19.2101 | -35.1655 | -29.5302
19.2162 | -35.166 | -29.5373
19.2223 | -35.1665 | -29.5444
19.2284 | -35.167 | -29.5516
19.2345 | -35.1674 | -29.5586
19.2406 | -35.1679 | -29.5654
19.2467 | -35.1684 | -29.5726
19.2528 | -35.1689 | -29.5797
19.2589 | -35.1693 | -29.5867
19.265 | -35.1698 | -29.5939
19.2712 | -35.1703 | -29.6009
19.2773 | -35.1708 | -29.6079
19.2834 | -35.1712 | -29.6149
19.2895 | -35.1717 | -29.6219
19.2956 | -35.1722 | -29.6291
19.3017 | -35.1727 | -29.6361
19.3078 | -35.1731 | -29.6431
19.3139 | -35.1736 | -29.6499
19.32 | -35.1741 | -29.6569
19.3261 | -35.1746 | -29.664
19.3322 | -35.175 | -29.671
19.3383 | -35.1755 | -29.678
19.3444 | -35.176 | -29.6849
19.3505 | -35.1765 | -29.692
19.3566 | -35.1769 | -29.699
19.3627 | -35.1774 | -29.7061
19.3688 | -35.1779 | -29.7131
19.3749 | -35.1784 | -29.7203
19.381 | -35.1788 | -29.727
19.3871 | -35.1793 | -29.7338
19.3932 | -35.1798 | -29.7409
19.3993 | -35.1803 | -29.748
19.4054 | -35.1807 | -29.7552
19.4115 | -35.1812 | -29.7623
19.4177 | -35.1817 | -29.7691
19.4238 | -35.1822 | -29.7761
19.4299 | -35.1826 | -29.783
19.436 | -35.1831 | -29.7902
19.4421 | -35.1836 | -29.7972
19.4482 | -35.1841 | -29.8043
19.4543 | -35.1845 | -29.811
19.4604 | -35.185 | -29.8181
19.4665 | -35.1855 | -29.825
19.4726 | -35.186 | -29.8318
19.4787 | -35.1864 | -29.8388
19.4848 | -35.1869 | -29.8458
19.4909 | -35.1874 | -29.8527
19.497 | -35.1879 | -29.8596
19.5031 | -35.1883 | -29.8665
19.5092 | -35.1888 | -29.8735
19.5153 | -35.1892 | -29.8805
19.5214 | -35.1897 | -29.8873
19.5275 | -35.1902 | -29.8943
19.5336 | -35.1907 | -29.9012
19.5397 | -35.1911 | -29.9081
19.5458 | -35.1916 | -29.9151
19.5519 | -35.1921 | -29.9221
19.5581 | -35.1925 | -29.929
19.5642 | -35.193 | -29.9358
19.5703 | -35.1935 | -29.9428
19.5764 | -35.194 | -29.9498
19.5825 | -35.1945 | -29.9566
19.5886 | -35.1949 | -29.9637
19.5947 | -35.1954 | -29.9703
19.6008 | -35.1958 | -29.9773
19.6069 | -35.1963 | -29.9843
19.613 | -35.1968 | -29.9913
19.6191 | -35.1972 | -29.9983
19.6252 | -35.1977 | -30.0052
19.6313 | -35.1982 | -30.0119
19.6374 | -35.1987 | -30.0188
19.6435 | -35.1992 | -30.0257
19.6496 | -35.1996 | -30.0327
19.6557 | -35.2001 | -30.0396
19.6618 | -35.2006 | -30.0463
19.6679 | -35.2011 | -30.0532
19.674 | -35.2015 | -30.0602
19.6801 | -35.202 | -30.0671
19.6862 | -35.2025 | -30.074
19.6923 | -35.2029 | -30.0808
19.6984 | -35.2034 | -30.0875
19.7046 | -35.2039 | -30.0945
19.7107 | -35.2044 | -30.1014
19.7168 | -35.2049 | -30.1083
19.7229 | -35.2053 | -30.1152
19.729 | -35.2058 | -30.1221
19.7351 | -35.2063 | -30.1288
19.7412 | -35.2068 | -30.1358
19.7473 | -35.2073 | -30.1428
19.7534 | -35.2077 | -30.1496
19.7595 | -35.2082 | -30.1566
19.7656 | -35.2087 | -30.1635
19.7717 | -35.2091 | -30.1702
19.7778 | -35.2096 | -30.1771
19.7839 | -35.2101 | -30.184
19.79 | -35.2105 | -30.1908
19.7961 | -35.211 | -30.1976
19.8022 | -35.2115 | -30.2044
19.8083 | -35.212 | -30.2113
19.8144 | -35.2124 | -30.2181
19.8205 | -35.2129 | -30.2249
19.8266 | -35.2134 | -30.2318
19.8327 | -35.2139 | -30.2387
19.8388 | -35.2143 | -30.2453
19.845 | -35.2148 | -30.2522
19.8511 | -35.2153 | -30.2591
19.8572 | -35.2157 | -30.266
19.8633 | -35.2162 | -30.2729
19.8694 | -35.2167 | -30.2799
19.8755 | -35.2172 | -30.2864
19.8816 | -35.2176 | -30.2934
19.8877 | -35.2181 | -30.3004
19.8938 | -35.2186 | -30.3071
19.8999 | -35.219 | -30.314
19.906 | -35.2195 | -30.3209
19.9121 | -35.22 | -30.3276
19.9182 | -35.2205 | -30.3345
19.9243 | -35.2209 | -30.3414
19.9304 | -35.2214 | -30.3483
19.9365 | -35.2218 | -30.3552
19.9426 | -35.2223 | -30.362
19.9487 | -35.2228 | -30.3687
19.9548 | -35.2233 | -30.3757
19.9609 | -35.2238 | -30.3826
19.967 | -35.2242 | -30.3895
19.9731 | -35.2247 | -30.3964
19.9792 | -35.2252 | -30.4033
19.9853 | -35.2256 | -30.4101
19.9915 | -35.2261 | -30.4168
19.9976 | -35.2266 | -30.4237
20.0037 | -35.227 | -30.4305
20.0098 | -35.2275 | -30.4374
20.0159 | -35.228 | -30.4441
20.022 | -35.2285 | -30.4508
20.0281 | -35.2289 | -30.4576
20.0342 | -35.2294 | -30.4646
20.0403 | -35.2299 | -30.4714
20.0464 | -35.2304 | -30.4783
20.0525 | -35.2308 | -30.4847
20.0586 | -35.2313 | -30.4918
20.0647 | -35.2318 | -30.4986
20.0708 | -35.2323 | -30.5056
20.0769 | -35.2327 | -30.5124
20.083 | -35.2332 | -30.5189
20.0891 | -35.2337 | -30.5258
20.0952 | -35.2341 | -30.5326
20.1013 | -35.2346 | -30.5394
20.1074 | -35.2351 | -30.5463
20.1135 | -35.2356 | -30.5531
20.1196 | -35.236 | -30.5597
20.1257 | -35.2365 | -30.5665
20.1319 | -35.237 | -30.5733
20.138 | -35.2374 | -30.5801
20.1441 | -35.2379 | -30.5869
20.1502 | -35.2384 | -30.5938
20.1563 | -35.2389 | -30.6003
20.1624 | -35.2394 | -30.6072
20.1685 | -35.2398 | -30.6141
20.1746 | -35.2403 | -30.6209
20.1807 | -35.2408 | -30.6278
20.1868 | -35.2413 | -30.6347
20.1929 | -35.2417 | -30.6411
20.199 | -35.2422 | -30.6478
20.2051 | -35.2427 | -30.6546
20.2112 | -35.2431 | -30.6614
20.2173 | -35.2436 | -30.6682
20.2234 | -35.2441 | -30.6749
20.2295 | -35.2446 | -30.6815
20.2356 | -35.245 | -30.6885
20.2417 | -35.2455 | -30.6952
20.2478 | -35.2459 | -30.7019
20.2539 | -35.2464 | -30.7087
20.26 | -35.2469 | -30.7152
20.2661 | -35.2473 | -30.7222
20.2723 | -35.2478 | -30.7289
20.2784 | -35.2483 | -30.7357
20.2845 | -35.2488 | -30.7423
20.2906 | -35.2492 | -30.7491
20.2967 | -35.2497 | -30.7557
20.3028 | -35.2502 | -30.7623
20.3089 | -35.2506 | -30.769
20.315 | -35.2511 | -30.7757
20.3211 | -35.2516 | -30.7825
20.3272 | -35.252 | -30.7894
20.3333 | -35.2525 | -30.7958
20.3394 | -35.253 | -30.8025
20.3455 | -35.2534 | -30.8091
20.3516 | -35.2539 | -30.8156
20.3577 | -35.2543 | -30.8224
20.3638 | -35.2548 | -30.829
20.3699 | -35.2553 | -30.8356
20.376 | -35.2557 | -30.8423
20.3821 | -35.2562 | -30.8492
20.3882 | -35.2567 | -30.856
20.3943 | -35.2572 | -30.8628
20.4004 | -35.2577 | -30.8693
20.4065 | -35.2581 | -30.8759
20.4126 | -35.2586 | -30.8827
20.4188 | -35.259 | -30.8897
20.4249 | -35.2595 | -30.8964
20.431 | -35.26 | -30.9031
20.4371 | -35.2605 | -30.9096
20.4432 | -35.2609 | -30.9164
20.4493 | -35.2614 | -30.9232
20.4554 | -35.2618 | -30.9299
20.4615 | -35.2623 | -30.9367
20.4676 | -35.2628 | -30.9436
20.4737 | -35.2632 | -30.95
20.4798 | -35.2637 | -30.9568
20.4859 | -35.2642 | -30.9636
20.492 | -35.2647 | -30.9702
20.4981 | -35.2651 | -30.9768
20.5042 | -35.2656 | -30.9835
20.5103 | -35.266 | -30.9899
20.5164 | -35.2665 | -30.9964
20.5225 | -35.2669 | -31.0031
20.5286 | -35.2674 | -31.0098
20.5347 | -35.2679 | -31.0163
20.5408 | -35.2684 | -31.0229
20.5469 | -35.2688 | -31.0298
20.553 | -35.2693 | -31.0367
20.5592 | -35.2697 | -31.0434
20.5653 | -35.2702 | -31.0502
20.5714 | -35.2707 | -31.0569
20.5775 | -35.2712 | -31.0637
20.5836 | -35.2716 | -31.0704
20.5897 | -35.2721 | -31.0773
20.5958 | -35.2725 | -31.0841
20.6019 | -35.273 | -31.0908
20.608 | -35.2734 | -31.0976
20.6141 | -35.2739 | -31.1042
20.6202 | -35.2744 | -31.1109
20.6263 | -35.2749 | -31.1175
20.6324 | -35.2753 | -31.1241
20.6385 | -35.2758 | -31.1309
20.6446 | -35.2763 | -31.1377
20.6507 | -35.2768 | -31.1438
20.6568 | -35.2772 | -31.1504
20.6629 | -35.2777 | -31.157
20.669 | -35.2782 | -31.1638
20.6751 | -35.2786 | -31.1704
20.6812 | -35.2791 | -31.1769
20.6873 | -35.2795 | -31.1836
20.6934 | -35.28 | -31.1903
20.6995 | -35.2805 | -31.197
20.7057 | -35.2809 | -31.2036
20.7118 | -35.2814 | -31.2104
20.7179 | -35.2819 | -31.2168
20.724 | -35.2823 | -31.2236
20.7301 | -35.2828 | -31.2301
20.7362 | -35.2833 | -31.2369
20.7423 | -35.2837 | -31.2435
20.7484 | -35.2842 | -31.2501
20.7545 | -35.2847 | -31.2564
20.7606 | -35.2851 | -31.2633
20.7667 | -35.2856 | -31.27
20.7728 | -35.2861 | -31.2766
20.7789 | -35.2865 | -31.2832
20.785 | -35.287 | -31.2899
20.7911 | -35.2875 | -31.2963
20.7972 | -35.2879 | -31.3031
20.8033 | -35.2884 | -31.3098
20.8094 | -35.2889 | -31.3163
20.8155 | -35.2893 | -31.3229
20.8216 | -35.2898 | -31.3293
20.8277 | -35.2902 | -31.3358
20.8338 | -35.2907 | -31.3426
20.8399 | -35.2912 | -31.3491
20.8461 | -35.2916 | -31.3558
20.8522 | -35.2921 | -31.3623
20.8583 | -35.2926 | -31.3687
20.8644 | -35.293 | -31.3753
20.8705 | -35.2935 | -31.3821
20.8766 | -35.294 | -31.3888
20.8827 | -35.2944 | -31.3953
20.8888 | -35.2949 | -31.4018
20.8949 | -35.2954 | -31.4082
20.901 | -35.2959 | -31.4149
20.9071 | -35.2963 | -31.4214
20.9132 | -35.2968 | -31.428
20.9193 | -35.2973 | -31.4347
20.9254 | -35.2978 | -31.4413
20.9315 | -35.2982 | -31.4478
20.9376 | -35.2987 | -31.4542
20.9437 | -35.2991 | -31.4607
20.9498 | -35.2996 | -31.4672
20.9559 | -35.3001 | -31.4739
20.962 | -35.3005 | -31.4804
20.9681 | -35.301 | -31.487
20.9742 | -35.3014 | -31.4936
20.9803 | -35.3019 | -31.5001
20.9864 | -35.3023 | -31.5067
20.9926 | -35.3028 | -31.5132
20.9987 | -35.3033 | -31.5197
21.0048 | -35.3037 | -31.5262
21.0109 | -35.3042 | -31.5329
21.017 | -35.3047 | -31.5395
21.0231 | -35.3051 | -31.546
21.0292 | -35.3056 | -31.5526
21.0353 | -35.3061 | -31.5591
21.0414 | -35.3066 | -31.5655
21.0475 | -35.307 | -31.5721
21.0536 | -35.3075 | -31.5785
21.0597 | -35.308 | -31.5851
21.0658 | -35.3085 | -31.5917
21.0719 | -35.3089 | -31.5982
21.078 | -35.3094 | -31.6049
21.0841 | -35.3099 | -31.6116
21.0902 | -35.3104 | -31.6176
21.0963 | -35.3109 | -31.6242
21.1024 | -35.3114 | -31.631
21.1085 | -35.3119 | -31.6373
21.1146 | -35.3124 | -31.6437
21.1207 | -35.3128 | -31.6502
21.1268 | -35.3133 | -31.6567
21.133 | -35.3137 | -31.6632
21.1391 | -35.3142 | -31.6697
21.1452 | -35.3147 | -31.6763
21.1513 | -35.3152 | -31.6828
21.1574 | -35.3156 | -31.6892
21.1635 | -35.3161 | -31.6956
21.1696 | -35.3166 | -31.7021
21.1757 | -35.317 | -31.7083
21.1818 | -35.3175 | -31.7151
21.1879 | -35.318 | -31.7214
21.194 | -35.3185 | -31.7278
21.2001 | -35.3189 | -31.7346
21.2062 | -35.3194 | -31.7413
21.2123 | -35.3199 | -31.7478
21.2184 | -35.3203 | -31.7542
21.2245 | -35.3208 | -31.7608
21.2306 | -35.3213 | -31.7671
21.2367 | -35.3217 | -31.7738
21.2428 | -35.3222 | -31.7804
21.2489 | -35.3227 | -31.7868
21.255 | -35.3231 | -31.7931
21.2611 | -35.3236 | -31.7999
21.2672 | -35.3241 | -31.8063
21.2733 | -35.3245 | -31.8129
21.2795 | -35.325 | -31.8192
21.2856 | -35.3254 | -31.8256
21.2917 | -35.3259 | -31.8321
21.2978 | -35.3264 | -31.8386
21.3039 | -35.3269 | -31.845
21.31 | -35.3273 | -31.8517
21.3161 | -35.3278 | -31.8583
21.3222 | -35.3282 | -31.8651
21.3283 | -35.3287 | -31.8716
21.3344 | -35.3291 | -31.8779
21.3405 | -35.3296 | -31.8844
21.3466 | -35.33 | -31.8906
21.3527 | -35.3305 | -31.897
21.3588 | -35.3309 | -31.9034
21.3649 | -35.3314 | -31.9099
21.371 | -35.3319 | -31.9162
21.3771 | -35.3323 | -31.9224
21.3832 | -35.3328 | -31.9287
21.3893 | -35.3333 | -31.9356
21.3954 | -35.3337 | -31.942
21.4015 | -35.3341 | -31.9486
21.4076 | -35.3346 | -31.9551
21.4137 | -35.3351 | -31.9615
21.4199 | -35.3356 | -31.9679
21.426 | -35.3361 | -31.9744
21.4321 | -35.3365 | -31.9808
21.4382 | -35.3369 | -31.9874
21.4443 | -35.3373 | -31.994
21.4504 | -35.3378 | -32.0003
21.4565 | -35.3382 | -32.0067
21.4626 | -35.3386 | -32.013
21.4687 | -35.3391 | -32.0195
21.4748 | -35.3396 | -32.0257
21.4809 | -35.3401 | -32.032
21.487 | -35.3405 | -32.0386
21.4931 | -35.3409 | -32.0448
21.4992 | -35.3414 | -32.051
21.5053 | -35.3419 | -32.0574
21.5114 | -35.3424 | -32.064
21.5175 | -35.3428 | -32.0705
21.5236 | -35.3432 | -32.0769
21.5297 | -35.3437 | -32.0835
21.5358 | -35.3442 | -32.0898
21.5419 | -35.3446 | -32.0966
21.548 | -35.3451 | -32.1033
21.5541 | -35.3455 | -32.1096
21.5602 | -35.346 | -32.1163
21.5664 | -35.3465 | -32.1227
21.5725 | -35.3469 | -32.1291
21.5786 | -35.3474 | -32.1356
21.5847 | -35.3479 | -32.1422
21.5908 | -35.3483 | -32.1485
21.5969 | -35.3488 | -32.155
21.603 | -35.3492 | -32.1616
21.6091 | -35.3497 | -32.168
21.6152 | -35.3502 | -32.1747
21.6213 | -35.3506 | -32.1812
21.6274 | -35.351 | -32.1875
21.6335 | -35.3515 | -32.1936
21.6396 | -35.352 | -32.2003
21.6457 | -35.3524 | -32.207
21.6518 | -35.3529 | -32.2135
21.6579 | -35.3534 | -32.2199
21.664 | -35.3538 | -32.2265
21.6701 | -35.3542 | -32.2329
21.6762 | -35.3547 | -32.2395
21.6823 | -35.3552 | -32.246
21.6884 | -35.3556 | -32.2527
21.6945 | -35.3561 | -32.2593
21.7006 | -35.3566 | -32.2659
21.7068 | -35.3571 | -32.2722
21.7129 | -35.3576 | -32.2784
21.719 | -35.358 | -32.285
21.7251 | -35.3584 | -32.2916
21.7312 | -35.3589 | -32.2982
21.7373 | -35.3594 | -32.3047
21.7434 | -35.3598 | -32.3112
21.7495 | -35.3603 | -32.3177
21.7556 | -35.3608 | -32.3241
21.7617 | -35.3613 | -32.3309
21.7678 | -35.3617 | -32.3375
21.7739 | -35.3622 | -32.3435
21.78 | -35.3627 | -32.3498
21.7861 | -35.3632 | -32.3566
21.7922 | -35.3636 | -32.3629
21.7983 | -35.364 | -32.3692
21.8044 | -35.3645 | -32.376
21.8105 | -35.3651 | -32.3819
21.8166 | -35.3655 | -32.3887
21.8227 | -35.366 | -32.3951
21.8288 | -35.3664 | -32.4018
21.8349 | -35.3669 | -32.4081
21.841 | -35.3673 | -32.4145
21.8471 | -35.3678 | -32.4213
21.8533 | -35.3683 | -32.4276
21.8594 | -35.3687 | -32.4338
21.8655 | -35.3692 | -32.44
21.8716 | -35.3696 | -32.4462
21.8777 | -35.3701 | -32.4522
21.8838 | -35.3706 | -32.4589
21.8899 | -35.371 | -32.4655
21.896 | -35.3716 | -32.472
21.9021 | -35.372 | -32.4785
21.9082 | -35.3725 | -32.4848
21.9143 | -35.3729 | -32.4913
21.9204 | -35.3733 | -32.4976
21.9265 | -35.3738 | -32.5041
21.9326 | -35.3743 | -32.5103
21.9387 | -35.3748 | -32.5169
21.9448 | -35.3752 | -32.5232
21.9509 | -35.3757 | -32.5295
21.957 | -35.3761 | -32.5361
21.9631 | -35.3766 | -32.5427
21.9692 | -35.377 | -32.549
21.9753 | -35.3775 | -32.5553
21.9814 | -35.3779 | -32.5617
21.9875 | -35.3784 | -32.5679
21.9937 | -35.3789 | -32.5744
21.9998 | -35.3793 | -32.5808
22.0059 | -35.3798 | -32.5874
22.012 | -35.3803 | -32.5938
22.0181 | -35.3808 | -32.6001
22.0242 | -35.3812 | -32.6066
22.0303 | -35.3817 | -32.6131
22.0364 | -35.3822 | -32.6193
22.0425 | -35.3826 | -32.6258
22.0486 | -35.3831 | -32.6323
22.0547 | -35.3836 | -32.6387
22.0608 | -35.384 | -32.6452
22.0669 | -35.3845 | -32.6516
22.073 | -35.3849 | -32.658
22.0791 | -35.3854 | -32.6645
22.0852 | -35.3858 | -32.6708
22.0913 | -35.3863 | -32.6774
22.0974 | -35.3868 | -32.6838
22.1035 | -35.3872 | -32.6904
22.1096 | -35.3876 | -32.6968
22.1157 | -35.3881 | -32.7031
22.1218 | -35.3886 | -32.7093
22.1279 | -35.3891 | -32.7158
22.134 | -35.3895 | -32.7223
22.1402 | -35.39 | -32.7288
22.1463 | -35.3905 | -32.7354
22.1524 | -35.391 | -32.7416
22.1585 | -35.3914 | -32.7479
22.1646 | -35.3919 | -32.7538
22.1707 | -35.3923 | -32.7601
22.1768 | -35.3928 | -32.7665
22.1829 | -35.3933 | -32.7729
22.189 | -35.3938 | -32.7792
22.1951 | -35.3942 | -32.7852
22.2012 | -35.3947 | -32.7918
22.2073 | -35.3951 | -32.7983
22.2134 | -35.3956 | -32.8047
22.2195 | -35.396 | -32.811
22.2256 | -35.3965 | -32.8174
22.2317 | -35.397 | -32.8237
22.2378 | -35.3974 | -32.8301
22.2439 | -35.3979 | -32.8363
22.25 | -35.3983 | -32.8426
22.2561 | -35.3988 | -32.8491
22.2622 | -35.3992 | -32.8555
22.2683 | -35.3997 | -32.8618
22.2744 | -35.4002 | -32.8681
22.2806 | -35.4006 | -32.8747
22.2867 | -35.4011 | -32.8811
22.2928 | -35.4015 | -32.8873
22.2989 | -35.402 | -32.894
22.305 | -35.4024 | -32.9003
22.3111 | -35.4029 | -32.9067
22.3172 | -35.4033 | -32.9132
22.3233 | -35.4038 | -32.9198
22.3294 | -35.4043 | -32.9259
22.3355 | -35.4047 | -32.932
22.3416 | -35.4052 | -32.9382
22.3477 | -35.4056 | -32.9447
22.3538 | -35.4061 | -32.951
22.3599 | -35.4065 | -32.9576
22.366 | -35.407 | -32.9639
22.3721 | -35.4075 | -32.9704
22.3782 | -35.4079 | -32.9769
22.3843 | -35.4084 | -32.9831
22.3904 | -35.4088 | -32.9893
22.3965 | -35.4093 | -32.9955
22.4026 | -35.4097 | -33.0017
22.4087 | -35.4102 | -33.0082
22.4148 | -35.4107 | -33.0145
22.4209 | -35.4112 | -33.0208
22.4271 | -35.4116 | -33.0271
22.4332 | -35.4121 | -33.0333
22.4393 | -35.4125 | -33.0396
22.4454 | -35.413 | -33.0455
22.4515 | -35.4135 | -33.0518
22.4576 | -35.414 | -33.058
22.4637 | -35.4144 | -33.0643
22.4698 | -35.4149 | -33.0705
22.4759 | -35.4153 | -33.0764
22.482 | -35.4158 | -33.083
22.4881 | -35.4162 | -33.0892
22.4942 | -35.4167 | -33.0954
22.5003 | -35.4172 | -33.1019
22.5064 | -35.4176 | -33.1083
22.5125 | -35.4181 | -33.1144
22.5186 | -35.4185 | -33.1209
22.5247 | -35.419 | -33.1273
22.5308 | -35.4194 | -33.1338
22.5369 | -35.4199 | -33.1401
22.543 | -35.4203 | -33.1465
22.5491 | -35.4208 | -33.1527
22.5552 | -35.4213 | -33.159
22.5613 | -35.4217 | -33.1656
22.5675 | -35.4222 | -33.1717
22.5736 | -35.4226 | -33.1782
22.5797 | -35.4231 | -33.1843
22.5858 | -35.4235 | -33.1901
22.5919 | -35.424 | -33.1965
22.598 | -35.4245 | -33.2029
22.6041 | -35.4249 | -33.2091
22.6102 | -35.4254 | -33.2152
22.6163 | -35.4258 | -33.2209
22.6224 | -35.4262 | -33.2272
22.6285 | -35.4267 | -33.2332
22.6346 | -35.4272 | -33.2394
22.6407 | -35.4276 | -33.2456
22.6468 | -35.4281 | -33.2519
22.6529 | -35.4286 | -33.2578
22.659 | -35.4291 | -33.2641
22.6651 | -35.4296 | -33.2703
22.6712 | -35.43 | -33.2764
22.6773 | -35.4305 | -33.2826
22.6834 | -35.431 | -33.2888
22.6895 | -35.4315 | -33.2949
22.6956 | -35.4319 | -33.3011
22.7017 | -35.4324 | -33.3076
22.7079 | -35.4329 | -33.3139
22.714 | -35.4333 | -33.3203
22.7201 | -35.4338 | -33.3268
22.7262 | -35.4342 | -33.3329
22.7323 | -35.4347 | -33.3392
22.7384 | -35.4351 | -33.3455
22.7445 | -35.4355 | -33.352
22.7506 | -35.436 | -33.3583
22.7567 | -35.4364 | -33.3646
22.7628 | -35.4369 | -33.3707
22.7689 | -35.4373 | -33.3769
22.775 | -35.4378 | -33.3832
22.7811 | -35.4383 | -33.3896
22.7872 | -35.4387 | -33.396
22.7933 | -35.4392 | -33.402
22.7994 | -35.4396 | -33.4083
22.8055 | -35.4401 | -33.4146
22.8116 | -35.4405 | -33.4208
22.8177 | -35.441 | -33.427
22.8238 | -35.4415 | -33.4334
22.8299 | -35.4419 | -33.4395
22.836 | -35.4424 | -33.4457
22.8421 | -35.4428 | -33.4519
22.8482 | -35.4432 | -33.4583
22.8544 | -35.4437 | -33.4646
22.8605 | -35.4442 | -33.4709
22.8666 | -35.4446 | -33.4772
22.8727 | -35.445 | -33.4836
22.8788 | -35.4455 | -33.4898
22.8849 | -35.4459 | -33.4959
22.891 | -35.4464 | -33.502
22.8971 | -35.4469 | -33.5083
22.9032 | -35.4474 | -33.5145
22.9093 | -35.4478 | -33.5209
22.9154 | -35.4482 | -33.5272
22.9215 | -35.4486 | -33.5336
22.9276 | -35.4491 | -33.5399
22.9337 | -35.4495 | -33.5459
22.9398 | -35.4499 | -33.552
22.9459 | -35.4504 | -33.5583
22.952 | -35.4509 | -33.5644
22.9581 | -35.4513 | -33.5705
22.9642 | -35.4518 | -33.5768
22.9703 | -35.4523 | -33.5829
22.9764 | -35.4528 | -33.5893
22.9825 | -35.4532 | -33.5954
22.9886 | -35.4537 | -33.6017
22.9948 | -35.4541 | -33.608
23.0009 | -35.4545 | -33.6142
23.007 | -35.455 | -33.6204
23.0131 | -35.4555 | -33.6268
23.0192 | -35.4559 | -33.633
23.0253 | -35.4563 | -33.6393
23.0314 | -35.4568 | -33.6457
23.0375 | -35.4573 | -33.6519
23.0436 | -35.4577 | -33.6583
23.0497 | -35.4582 | -33.6646
23.0558 | -35.4586 | -33.671
23.0619 | -35.4591 | -33.6772
23.068 | -35.4595 | -33.6834
23.0741 | -35.46 | -33.6894
23.0802 | -35.4604 | -33.6956
23.0863 | -35.4609 | -33.7019
23.0924 | -35.4613 | -33.708
23.0985 | -35.4618 | -33.7142
23.1046 | -35.4622 | -33.7205
23.1107 | -35.4626 | -33.7265
23.1168 | -35.4631 | -33.7328
23.1229 | -35.4636 | -33.739
23.129 | -35.464 | -33.7453
23.1351 | -35.4644 | -33.7513
23.1413 | -35.4649 | -33.7575
23.1474 | -35.4654 | -33.7636
23.1535 | -35.4658 | -33.7699
23.1596 | -35.4663 | -33.7758
23.1657 | -35.4667 | -33.7822
23.1718 | -35.4672 | -33.7883
23.1779 | -35.4676 | -33.7946
23.184 | -35.4681 | -33.8005
23.1901 | -35.4686 | -33.8069
23.1962 | -35.469 | -33.8131
23.2023 | -35.4695 | -33.8195
23.2084 | -35.4699 | -33.8256
23.2145 | -35.4704 | -33.8315
23.2206 | -35.4708 | -33.8378
23.2267 | -35.4713 | -33.8439
23.2328 | -35.4718 | -33.85
23.2389 | -35.4722 | -33.8562
23.245 | -35.4727 | -33.8624
23.2511 | -35.4731 | -33.8683
23.2572 | -35.4736 | -33.8744
23.2633 | -35.474 | -33.8807
23.2694 | -35.4745 | -33.8869
23.2755 | -35.4749 | -33.8932
23.2817 | -35.4754 | -33.8993
23.2878 | -35.4758 | -33.9053
23.2939 | -35.4763 | -33.9114
23.3 | -35.4767 | -33.9178
23.3061 | -35.4772 | -33.9242
23.3122 | -35.4776 | -33.9303
23.3183 | -35.4781 | -33.9363
23.3244 | -35.4785 | -33.9425
23.3305 | -35.479 | -33.9485
23.3366 | -35.4794 | -33.9548
23.3427 | -35.4799 | -33.9607
23.3488 | -35.4803 | -33.9671
23.3549 | -35.4808 | -33.9733
23.361 | -35.4813 | -33.9795
23.3671 | -35.4817 | -33.9855
23.3732 | -35.4822 | -33.9915
23.3793 | -35.4827 | -33.9976
23.3854 | -35.4831 | -34.0039
23.3915 | -35.4836 | -34.0102
23.3976 | -35.484 | -34.0162
23.4037 | -35.4845 | -34.0223
23.4098 | -35.4849 | -34.0284
23.4159 | -35.4854 | -34.0347
23.422 | -35.4859 | -34.041
23.4282 | -35.4863 | -34.0471
23.4343 | -35.4868 | -34.0533
23.4404 | -35.4872 | -34.0596
23.4465 | -35.4877 | -34.0655
23.4526 | -35.4881 | -34.0717
23.4587 | -35.4886 | -34.0778
23.4648 | -35.4891 | -34.0837
23.4709 | -35.4895 | -34.0898
23.477 | -35.4899 | -34.0957
23.4831 | -35.4904 | -34.1019
23.4892 | -35.4908 | -34.108
23.4953 | -35.4913 | -34.1142
23.5014 | -35.4917 | -34.1205
23.5075 | -35.4922 | -34.1268
23.5136 | -35.4926 | -34.1328
23.5197 | -35.4931 | -34.139
23.5258 | -35.4935 | -34.1451
23.5319 | -35.494 | -34.1512
23.538 | -35.4944 | -34.1575
23.5441 | -35.4949 | -34.1635
23.5502 | -35.4953 | -34.1695
23.5563 | -35.4958 | -34.1755
23.5624 | -35.4962 | -34.1817
23.5686 | -35.4966 | -34.1877
23.5747 | -35.4971 | -34.1937
23.5808 | -35.4975 | -34.1997
23.5869 | -35.498 | -34.2058
23.593 | -35.4984 | -34.2121
23.5991 | -35.4989 | -34.2184
23.6052 | -35.4993 | -34.2242
23.6113 | -35.4998 | -34.2305
23.6174 | -35.5002 | -34.2366
23.6235 | -35.5006 | -34.2428
23.6296 | -35.5011 | -34.2488
23.6357 | -35.5015 | -34.2551
23.6418 | -35.502 | -34.2612
23.6479 | -35.5024 | -34.2674
23.654 | -35.5029 | -34.2736
23.6601 | -35.5033 | -34.2799
23.6662 | -35.5037 | -34.2858
23.6723 | -35.5042 | -34.2919
23.6784 | -35.5046 | -34.2979
23.6845 | -35.5051 | -34.3039
23.6906 | -35.5055 | -34.3101
23.6967 | -35.506 | -34.3162
23.7028 | -35.5064 | -34.3225
23.7089 | -35.5068 | -34.3284
23.7151 | -35.5073 | -34.3345
23.7212 | -35.5077 | -34.3408
23.7273 | -35.5082 | -34.3469
23.7334 | -35.5086 | -34.3528
23.7395 | -35.509 | -34.359
23.7456 | -35.5095 | -34.3649
23.7517 | -35.5099 | -34.3708
23.7578 | -35.5103 | -34.377
23.7639 | -35.5108 | -34.383
23.77 | -35.5112 | -34.3892
23.7761 | -35.5116 | -34.3952
23.7822 | -35.5121 | -34.4011
23.7883 | -35.5125 | -34.4072
23.7944 | -35.513 | -34.4134
23.8005 | -35.5134 | -34.4195
23.8066 | -35.5139 | -34.4256
23.8127 | -35.5143 | -34.4315
23.8188 | -35.5147 | -34.4378
23.8249 | -35.5152 | -34.4439
23.831 | -35.5156 | -34.45
23.8371 | -35.516 | -34.4559
23.8432 | -35.5165 | -34.4621
23.8493 | -35.5169 | -34.468
23.8555 | -35.5174 | -34.4741
23.8616 | -35.5178 | -34.4801
23.8677 | -35.5183 | -34.486
23.8738 | -35.5187 | -34.4921
23.8799 | -35.5191 | -34.4983
23.886 | -35.5196 | -34.5041
23.8921 | -35.52 | -34.5102
23.8982 | -35.5204 | -34.5163
23.9043 | -35.5209 | -34.5224
23.9104 | -35.5213 | -34.5285
23.9165 | -35.5217 | -34.5345
23.9226 | -35.5222 | -34.5404
23.9287 | -35.5226 | -34.5465
23.9348 | -35.523 | -34.5525
23.9409 | -35.5234 | -34.5585
23.947 | -35.5239 | -34.5645
23.9531 | -35.5243 | -34.5702
23.9592 | -35.5247 | -34.5764
23.9653 | -35.5252 | -34.5825
23.9714 | -35.5256 | -34.5886
23.9775 | -35.5261 | -34.5948
23.9836 | -35.5265 | -34.6009
23.9897 | -35.5269 | -34.6069
23.9958 | -35.5274 | -34.6129
24.002 | -35.5278 | -34.619
24.0081 | -35.5282 | -34.6249
24.0142 | -35.5286 | -34.6309
24.0203 | -35.529 | -34.637
24.0264 | -35.5295 | -34.6431
24.0325 | -35.5299 | -34.649
24.0386 | -35.5303 | -34.655
24.0447 | -35.5307 | -34.6613
24.0508 | -35.5312 | -34.6672
24.0569 | -35.5316 | -34.6731
24.063 | -35.5321 | -34.6789
24.0691 | -35.5325 | -34.6849
24.0752 | -35.5329 | -34.6912
24.0813 | -35.5334 | -34.6973
24.0874 | -35.5338 | -34.7031
24.0935 | -35.5343 | -34.7093
24.0996 | -35.5347 | -34.7152
24.1057 | -35.5351 | -34.7212
24.1118 | -35.5356 | -34.7273
24.1179 | -35.536 | -34.7335
24.124 | -35.5365 | -34.7395
24.1301 | -35.5369 | -34.7451
24.1362 | -35.5373 | -34.7511
24.1424 | -35.5378 | -34.757
24.1485 | -35.5382 | -34.7631
24.1546 | -35.5386 | -34.769
24.1607 | -35.5391 | -34.7751
24.1668 | -35.5395 | -34.7811
24.1729 | -35.5399 | -34.7873
24.179 | -35.5403 | -34.7933
24.1851 | -35.5407 | -34.7994
24.1912 | -35.5412 | -34.8052
24.1973 | -35.5416 | -34.8114
24.2034 | -35.542 | -34.8175
24.2095 | -35.5425 | -34.8236
24.2156 | -35.5429 | -34.8298
24.2217 | -35.5433 | -34.8359
24.2278 | -35.5437 | -34.8419
24.2339 | -35.5441 | -34.848
24.24 | -35.5446 | -34.8538
24.2461 | -35.545 | -34.8598
24.2522 | -35.5454 | -34.8659
24.2583 | -35.5458 | -34.8719
24.2644 | -35.5463 | -34.878
24.2705 | -35.5467 | -34.884
24.2766 | -35.5471 | -34.8901
24.2827 | -35.5475 | -34.896
24.2889 | -35.548 | -34.902
24.295 | -35.5484 | -34.908
24.3011 | -35.5488 | -34.9139
24.3072 | -35.5492 | -34.9196
24.3133 | -35.5496 | -34.9256
24.3194 | -35.5501 | -34.9316
24.3255 | -35.5505 | -34.9378
24.3316 | -35.5509 | -34.9437
24.3377 | -35.5513 | -34.9497
24.3438 | -35.5518 | -34.9554
24.3499 | -35.5522 | -34.9615
24.356 | -35.5526 | -34.9675
24.3621 | -35.553 | -34.9735
24.3682 | -35.5534 | -34.9794
24.3743 | -35.5539 | -34.9855
24.3804 | -35.5543 | -34.9914
24.3865 | -35.5547 | -34.9973
24.3926 | -35.5551 | -35.0034
24.3987 | -35.5555 | -35.0095
24.4048 | -35.556 | -35.0155
24.4109 | -35.5563 | -35.0213
24.417 | -35.5567 | -35.0273
24.4231 | -35.5572 | -35.0333
24.4293 | -35.5576 | -35.0393
24.4354 | -35.558 | -35.0453
24.4415 | -35.5585 | -35.0511
24.4476 | -35.5589 | -35.0568
24.4537 | -35.5593 | -35.0628
24.4598 | -35.5597 | -35.0688
24.4659 | -35.5601 | -35.0747
24.472 | -35.5606 | -35.0807
24.4781 | -35.561 | -35.0866
24.4842 | -35.5614 | -35.0925
24.4903 | -35.5618 | -35.0985
24.4964 | -35.5622 | -35.1045
24.5025 | -35.5626 | -35.1103
24.5086 | -35.5631 | -35.1163
24.5147 | -35.5635 | -35.1223
24.5208 | -35.5639 | -35.128
24.5269 | -35.5643 | -35.1338
24.533 | -35.5647 | -35.1397
24.5391 | -35.5651 | -35.1456
24.5452 | -35.5656 | -35.1515
24.5513 | -35.5659 | -35.1575
24.5574 | -35.5664 | -35.1633
24.5635 | -35.5668 | -35.1691
24.5696 | -35.5672 | -35.175
24.5758 | -35.5676 | -35.1808
24.5819 | -35.568 | -35.1866
24.588 | -35.5684 | -35.1923
24.5941 | -35.5688 | -35.1981
24.6002 | -35.5692 | -35.204
24.6063 | -35.5697 | -35.2098
24.6124 | -35.5701 | -35.2156
24.6185 | -35.5705 | -35.2216
24.6246 | -35.5709 | -35.2274
24.6307 | -35.5713 | -35.2332
24.6368 | -35.5717 | -35.2393
24.6429 | -35.5722 | -35.2454
24.649 | -35.5726 | -35.2512
24.6551 | -35.573 | -35.257
24.6612 | -35.5734 | -35.2625
24.6673 | -35.5739 | -35.2687
24.6734 | -35.5743 | -35.2745
24.6795 | -35.5747 | -35.2804
24.6856 | -35.5751 | -35.2863
24.6917 | -35.5755 | -35.292
24.6978 | -35.5759 | -35.2979
24.7039 | -35.5763 | -35.3039
24.71 | -35.5767 | -35.3098
24.7162 | -35.5771 | -35.3156
24.7223 | -35.5775 | -35.3214
24.7284 | -35.5779 | -35.3272
24.7345 | -35.5783 | -35.3329
24.7406 | -35.5787 | -35.3389
24.7467 | -35.5791 | -35.3449
24.7528 | -35.5796 | -35.3508
24.7589 | -35.58 | -35.3567
24.765 | -35.5804 | -35.3622
24.7711 | -35.5808 | -35.3682
24.7772 | -35.5812 | -35.3741
24.7833 | -35.5816 | -35.3801
24.7894 | -35.582 | -35.386
24.7955 | -35.5824 | -35.392
24.8016 | -35.5828 | -35.3976
24.8077 | -35.5832 | -35.4036
24.8138 | -35.5837 | -35.4095
24.8199 | -35.5841 | -35.4152
24.826 | -35.5845 | -35.421
24.8321 | -35.5849 | -35.427
24.8382 | -35.5852 | -35.4325
24.8443 | -35.5857 | -35.4382
24.8504 | -35.5861 | -35.444
24.8565 | -35.5865 | -35.4497
24.8627 | -35.5869 | -35.4554
24.8688 | -35.5872 | -35.4611
24.8749 | -35.5877 | -35.4671
24.881 | -35.5881 | -35.473
24.8871 | -35.5885 | -35.479
24.8932 | -35.5889 | -35.485
24.8993 | -35.5893 | -35.491
24.9054 | -35.5897 | -35.4966
24.9115 | -35.5901 | -35.5023
24.9176 | -35.5905 | -35.5084
24.9237 | -35.5909 | -35.5141
24.9298 | -35.5913 | -35.5199
24.9359 | -35.5917 | -35.5258
24.942 | -35.5921 | -35.5316
24.9481 | -35.5926 | -35.5374
24.9542 | -35.5929 | -35.5433
24.9603 | -35.5933 | -35.5494
24.9664 | -35.5938 | -35.5553
24.9725 | -35.5942 | -35.5613
24.9786 | -35.5946 | -35.5669
24.9847 | -35.595 | -35.5728
24.9908 | -35.5954 | -35.5787
24.9969 | -35.5958 | -35.5845
25.0031 | -35.5962 | -35.5903
25.0092 | -35.5966 | -35.5959
25.0153 | -35.597 | -35.6017
25.0214 | -35.5974 | -35.6075
25.0275 | -35.5978 | -35.6131
25.0336 | -35.5982 | -35.6191
25.0397 | -35.5987 | -35.625
25.0458 | -35.599 | -35.6308
25.0519 | -35.5995 | -35.6365
25.058 | -35.5998 | -35.6424
25.0641 | -35.6003 | -35.6482
25.0702 | -35.6006 | -35.6541
25.0763 | -35.6011 | -35.66
25.0824 | -35.6015 | -35.6655
25.0885 | -35.6019 | -35.6713
25.0946 | -35.6022 | -35.6772
25.1007 | -35.6026 | -35.6831
25.1068 | -35.603 | -35.6891
25.1129 | -35.6034 | -35.695
25.119 | -35.6038 | -35.7004
25.1251 | -35.6042 | -35.7061
25.1312 | -35.6046 | -35.712
25.1373 | -35.605 | -35.7178
25.1435 | -35.6055 | -35.7238
25.1496 | -35.6058 | -35.7295
25.1557 | -35.6062 | -35.7354
25.1618 | -35.6066 | -35.7413
25.1679 | -35.607 | -35.7472
25.174 | -35.6074 | -35.753
25.1801 | -35.6078 | -35.7589
25.1862 | -35.6082 | -35.7646
25.1923 | -35.6086 | -35.7706
25.1984 | -35.609 | -35.7764
25.2045 | -35.6094 | -35.7823
25.2106 | -35.6098 | -35.7883
25.2167 | -35.6101 | -35.7941
25.2228 | -35.6105 | -35.7997
25.2289 | -35.6109 | -35.8056
25.235 | -35.6113 | -35.8113
25.2411 | -35.6118 | -35.8172
25.2472 | -35.6122 | -35.823
25.2533 | -35.6126 | -35.8289
25.2594 | -35.6129 | -35.8345
25.2655 | -35.6133 | -35.8405
25.2716 | -35.6137 | -35.8462
25.2777 | -35.6141 | -35.8518
25.2838 | -35.6146 | -35.8578
25.29 | -35.615 | -35.8637
25.2961 | -35.6154 | -35.8691
25.3022 | -35.6158 | -35.8749
25.3083 | -35.6162 | -35.8805
25.3144 | -35.6166 | -35.8863
25.3205 | -35.617 | -35.8923
25.3266 | -35.6174 | -35.8979
25.3327 | -35.6178 | -35.904
25.3388 | -35.6182 | -35.9097
25.3449 | -35.6186 | -35.9156
25.351 | -35.619 | -35.9213
25.3571 | -35.6194 | -35.9272
25.3632 | -35.6198 | -35.9329
25.3693 | -35.6202 | -35.9389
25.3754 | -35.6206 | -35.9447
25.3815 | -35.6209 | -35.9506
25.3876 | -35.6213 | -35.9564
25.3937 | -35.6217 | -35.9622
25.3998 | -35.6221 | -35.9678
25.4059 | -35.6224 | -35.9734
25.412 | -35.6228 | -35.9793
25.4181 | -35.6232 | -35.9851
25.4242 | -35.6236 | -35.9908
25.4304 | -35.624 | -35.9968
25.4365 | -35.6243 | -36.0025
25.4426 | -35.6247 | -36.0084
25.4487 | -35.625 | -36.0145
25.4548 | -35.6254 | -36.0201
25.4609 | -35.6258 | -36.0257
25.467 | -35.6262 | -36.031
25.4731 | -35.6266 | -36.0365
25.4792 | -35.627 | -36.0424
25.4853 | -35.6273 | -36.0482
25.4914 | -35.6277 | -36.0537
25.4975 | -35.6281 | -36.0593
25.5036 | -35.6285 | -36.0647
25.5097 | -35.6289 | -36.0706
25.5158 | -35.6293 | -36.0762
25.5219 | -35.6297 | -36.0822
25.528 | -35.6301 | -36.0878
25.5341 | -35.6305 | -36.0933
25.5402 | -35.6309 | -36.0988
25.5463 | -35.6312 | -36.1047
25.5524 | -35.6316 | -36.1101
25.5585 | -35.6319 | -36.1157
25.5646 | -35.6323 | -36.1217
25.5707 | -35.6327 | -36.1273
25.5769 | -35.6331 | -36.1326
25.583 | -35.6334 | -36.1382
25.5891 | -35.6338 | -36.1437
25.5952 | -35.6342 | -36.1494
25.6013 | -35.6345 | -36.1552
25.6074 | -35.635 | -36.1609
25.6135 | -35.6353 | -36.1666
25.6196 | -35.6357 | -36.1727
25.6257 | -35.636 | -36.1785
25.6318 | -35.6364 | -36.1842
25.6379 | -35.6367 | -36.1898
25.644 | -35.637 | -36.1956
25.6501 | -35.6374 | -36.2013
25.6562 | -35.6378 | -36.2071
25.6623 | -35.6382 | -36.2126
25.6684 | -35.6385 | -36.2185
25.6745 | -35.6389 | -36.2245
25.6806 | -35.6393 | -36.2304
25.6867 | -35.6396 | -36.2362
25.6928 | -35.64 | -36.2418
25.6989 | -35.6404 | -36.2476
25.705 | -35.6408 | -36.2535
25.7111 | -35.6412 | -36.2592
25.7173 | -35.6416 | -36.265
25.7234 | -35.642 | -36.2705
25.7295 | -35.6424 | -36.276
25.7356 | -35.6427 | -36.2819
25.7417 | -35.6431 | -36.2877
25.7478 | -35.6434 | -36.2933
25.7539 | -35.6438 | -36.2989
25.76 | -35.6441 | -36.3046
25.7661 | -35.6445 | -36.3105
25.7722 | -35.6449 | -36.3162
25.7783 | -35.6453 | -36.3222
25.7844 | -35.6456 | -36.3279
25.7905 | -35.646 | -36.3335
25.7966 | -35.6463 | -36.3392
25.8027 | -35.6466 | -36.3449
25.8088 | -35.647 | -36.3506
25.8149 | -35.6474 | -36.3564
25.821 | -35.6477 | -36.3618
25.8271 | -35.6481 | -36.3676
25.8332 | -35.6485 | -36.3733
25.8393 | -35.6489 | -36.3792
25.8454 | -35.6492 | -36.385
25.8515 | -35.6496 | -36.3905
25.8576 | -35.6499 | -36.3958
25.8638 | -35.6504 | -36.4015
25.8699 | -35.6507 | -36.4072
25.876 | -35.651 | -36.4127
25.8821 | -35.6514 | -36.4183
25.8882 | -35.6518 | -36.424
25.8943 | -35.6522 | -36.4296
25.9004 | -35.6526 | -36.4354
25.9065 | -35.6529 | -36.441
25.9126 | -35.6533 | -36.4466
25.9187 | -35.6537 | -36.4522
25.9248 | -35.654 | -36.4575
25.9309 | -35.6544 | -36.4632
25.937 | -35.6548 | -36.4687
25.9431 | -35.6551 | -36.4747
25.9492 | -35.6555 | -36.4804
25.9553 | -35.6559 | -36.486
25.9614 | -35.6562 | -36.4915
25.9675 | -35.6566 | -36.497
25.9736 | -35.657 | -36.5026
25.9797 | -35.6573 | -36.5084
25.9858 | -35.6577 | -36.514
25.9919 | -35.6581 | -36.5197
25.998 | -35.6585 | -36.525
26.0042 | -35.6588 | -36.5307
26.0103 | -35.6592 | -36.5362
26.0164 | -35.6596 | -36.5417
26.0225 | -35.6599 | -36.5474
26.0286 | -35.6603 | -36.5528
26.0347 | -35.6606 | -36.5583
26.0408 | -35.6609 | -36.5638
26.0469 | -35.6613 | -36.5696
26.053 | -35.6616 | -36.5753
26.0591 | -35.662 | -36.5811
26.0652 | -35.6623 | -36.5864
26.0713 | -35.6627 | -36.5921
26.0774 | -35.6631 | -36.5979
26.0835 | -35.6634 | -36.6038
26.0896 | -35.6638 | -36.6095
26.0957 | -35.6642 | -36.615
26.1018 | -35.6645 | -36.6201
26.1079 | -35.6649 | -36.6256
26.114 | -35.6652 | -36.631
26.1201 | -35.6656 | -36.6369
26.1262 | -35.6659 | -36.6427
26.1323 | -35.6663 | -36.6484
26.1384 | -35.6667 | -36.6537
26.1445 | -35.6671 | -36.6592
26.1507 | -35.6674 | -36.6651
26.1568 | -35.6678 | -36.6709
26.1629 | -35.6681 | -36.6766
26.169 | -35.6685 | -36.682
26.1751 | -35.6689 | -36.6872
26.1812 | -35.6692 | -36.693
26.1873 | -35.6696 | -36.6988
26.1934 | -35.6699 | -36.7042
26.1995 | -35.6703 | -36.7099
26.2056 | -35.6707 | -36.7159
26.2117 | -35.671 | -36.7214
26.2178 | -35.6714 | -36.7268
26.2239 | -35.6717 | -36.7323
26.23 | -35.6721 | -36.7377
26.2361 | -35.6724 | -36.7436
26.2422 | -35.6728 | -36.7487
26.2483 | -35.6731 | -36.754
26.2544 | -35.6735 | -36.7594
26.2605 | -35.6738 | -36.7648
26.2666 | -35.6741 | -36.7707
26.2727 | -35.6745 | -36.7762
26.2788 | -35.6748 | -36.7818
26.2849 | -35.6751 | -36.7873
26.2911 | -35.6755 | -36.7928
26.2972 | -35.6759 | -36.7986
26.3033 | -35.6762 | -36.8043
26.3094 | -35.6766 | -36.8095
26.3155 | -35.6769 | -36.8149
26.3216 | -35.6772 | -36.8207
26.3277 | -35.6776 | -36.8261
26.3338 | -35.6779 | -36.8316
26.3399 | -35.6783 | -36.8372
26.346 | -35.6786 | -36.8427
26.3521 | -35.679 | -36.8482
26.3582 | -35.6793 | -36.8537
26.3643 | -35.6797 | -36.8591
26.3704 | -35.68 | -36.8648
26.3765 | -35.6803 | -36.8704
26.3826 | -35.6807 | -36.876
26.3887 | -35.6811 | -36.8817
26.3948 | -35.6814 | -36.8873
26.4009 | -35.6818 | -36.8927
26.407 | -35.6821 | -36.8982
26.4131 | -35.6824 | -36.9037
26.4192 | -35.6828 | -36.9086
26.4253 | -35.6831 | -36.9139
26.4314 | -35.6835 | -36.9192
26.4376 | -35.6838 | -36.9249
26.4437 | -35.6842 | -36.9307
26.4498 | -35.6846 | -36.9366
26.4559 | -35.6849 | -36.9421
26.462 | -35.6853 | -36.9476
26.4681 | -35.6856 | -36.953
26.4742 | -35.686 | -36.9586
26.4803 | -35.6863 | -36.9639
26.4864 | -35.6867 | -36.9689
26.4925 | -35.687 | -36.9744
26.4986 | -35.6874 | -36.9797
26.5047 | -35.6878 | -36.9852
26.5108 | -35.6881 | -36.9906
26.5169 | -35.6885 | -36.9961
26.523 | -35.6889 | -37.0016
26.5291 | -35.6893 | -37.0069
26.5352 | -35.6896 | -37.0125
26.5413 | -35.69 | -37.0183
26.5474 | -35.6904 | -37.0239
26.5535 | -35.6907 | -37.0296
26.5596 | -35.691 | -37.0349
26.5657 | -35.6914 | -37.0405
26.5718 | -35.6917 | -37.0458
26.578 | -35.6921 | -37.0514
26.5841 | -35.6924 | -37.0571
26.5902 | -35.6927 | -37.0629
26.5963 | -35.693 | -37.068
26.6024 | -35.6934 | -37.0736
26.6085 | -35.6937 | -37.0791
26.6146 | -35.6941 | -37.0848
26.6207 | -35.6944 | -37.0904
26.6268 | -35.6948 | -37.0959
26.6329 | -35.6951 | -37.1013
26.639 | -35.6954 | -37.1068
26.6451 | -35.6958 | -37.1125
26.6512 | -35.6961 | -37.1182
26.6573 | -35.6964 | -37.1237
26.6634 | -35.6967 | -37.1289
26.6695 | -35.6971 | -37.1347
26.6756 | -35.6974 | -37.1402
26.6817 | -35.6977 | -37.1459
26.6878 | -35.6981 | -37.1515
26.6939 | -35.6985 | -37.1569
26.7 | -35.6988 | -37.1625
26.7061 | -35.6991 | -37.168
26.7122 | -35.6995 | -37.1737
26.7183 | -35.6999 | -37.1793
26.7245 | -35.7002 | -37.1851
26.7306 | -35.7005 | -37.1904
26.7367 | -35.7008 | -37.1959
26.7428 | -35.7012 | -37.2013
26.7489 | -35.7015 | -37.2069
26.755 | -35.7018 | -37.2122
26.7611 | -35.7021 | -37.2175
26.7672 | -35.7025 | -37.2231
26.7733 | -35.7028 | -37.2289
26.7794 | -35.7032 | -37.2343
26.7855 | -35.7036 | -37.2399
26.7916 | -35.7039 | -37.2455
26.7977 | -35.7043 | -37.2508
26.8038 | -35.7046 | -37.2562
26.8099 | -35.705 | -37.2614
26.816 | -35.7053 | -37.2666
26.8221 | -35.7056 | -37.2722
26.8282 | -35.706 | -37.2775
26.8343 | -35.7063 | -37.2832
26.8404 | -35.7067 | -37.2885
26.8465 | -35.707 | -37.2938
26.8526 | -35.7074 | -37.2991
26.8587 | -35.7077 | -37.3043
26.8649 | -35.708 | -37.3096
26.871 | -35.7084 | -37.315
26.8771 | -35.7088 | -37.3205
26.8832 | -35.7091 | -37.3258
26.8893 | -35.7095 | -37.3313
26.8954 | -35.7099 | -37.3364
26.9015 | -35.7102 | -37.3416
26.9076 | -35.7105 | -37.3468
26.9137 | -35.7109 | -37.3519
26.9198 | -35.7112 | -37.3573
26.9259 | -35.7116 | -37.3629
26.932 | -35.7119 | -37.3684
26.9381 | -35.7122 | -37.3736
26.9442 | -35.7125 | -37.3788
26.9503 | -35.7129 | -37.384
26.9564 | -35.7133 | -37.3893
26.9625 | -35.7136 | -37.3946
26.9686 | -35.7139 | -37.4002
26.9747 | -35.7142 | -37.4055
26.9808 | -35.7146 | -37.4106
26.9869 | -35.7149 | -37.416
26.993 | -35.7153 | -37.4212
26.9991 | -35.7156 | -37.4265
27.0052 | -35.716 | -37.4319
27.0114 | -35.7163 | -37.4375
27.0175 | -35.7166 | -37.4424
27.0236 | -35.7169 | -37.4479
27.0297 | -35.7173 | -37.453
27.0358 | -35.7176 | -37.4584
27.0419 | -35.718 | -37.4641
27.048 | -35.7183 | -37.4697
27.0541 | -35.7187 | -37.4749
27.0602 | -35.719 | -37.48
27.0663 | -35.7193 | -37.4857
27.0724 | -35.7196 | -37.4912
27.0785 | -35.72 | -37.4969
27.0846 | -35.7203 | -37.5024
27.0907 | -35.7207 | -37.5077
27.0968 | -35.721 | -37.5128
27.1029 | -35.7213 | -37.5181
27.109 | -35.7217 | -37.5236
27.1151 | -35.722 | -37.5293
27.1212 | -35.7223 | -37.5346
27.1273 | -35.7226 | -37.5401
27.1334 | -35.723 | -37.5455
27.1395 | -35.7233 | -37.5506
27.1456 | -35.7236 | -37.5559
27.1518 | -35.724 | -37.5613
27.1579 | -35.7243 | -37.566
27.164 | -35.7247 | -37.5713
27.1701 | -35.725 | -37.5769
27.1762 | -35.7253 | -37.5821
27.1823 | -35.7256 | -37.5873
27.1884 | -35.7259 | -37.5926
27.1945 | -35.7263 | -37.5974
27.2006 | -35.7266 | -37.603
27.2067 | -35.7269 | -37.6083
27.2128 | -35.7272 | -37.6136
27.2189 | -35.7275 | -37.619
27.225 | -35.7279 | -37.6245
27.2311 | -35.7282 | -37.6298
27.2372 | -35.7285 | -37.6353
27.2433 | -35.7288 | -37.6406
27.2494 | -35.7292 | -37.646
27.2555 | -35.7295 | -37.6513
27.2616 | -35.7299 | -37.6563
27.2677 | -35.7302 | -37.6616
27.2738 | -35.7305 | -37.6668
27.2799 | -35.7308 | -37.6722
27.286 | -35.7312 | -37.6775
27.2921 | -35.7315 | -37.6829
27.2983 | -35.7318 | -37.6879
27.3044 | -35.7321 | -37.6934
27.3105 | -35.7324 | -37.6988
27.3166 | -35.7327 | -37.7039
27.3227 | -35.7331 | -37.7091
27.3288 | -35.7334 | -37.7146
27.3349 | -35.7337 | -37.7193
27.341 | -35.7341 | -37.7247
27.3471 | -35.7344 | -37.7298
27.3532 | -35.7347 | -37.7352
27.3593 | -35.7351 | -37.7405
27.3654 | -35.7354 | -37.7459
27.3715 | -35.7357 | -37.7509
27.3776 | -35.736 | -37.7564
27.3837 | -35.7363 | -37.7616
27.3898 | -35.7366 | -37.7669
27.3959 | -35.737 | -37.7722
27.402 | -35.7373 | -37.777
27.4081 | -35.7376 | -37.7824
27.4142 | -35.7379 | -37.7878
27.4203 | -35.7383 | -37.7931
27.4264 | -35.7386 | -37.7984
27.4325 | -35.739 | -37.8035
27.4387 | -35.7393 | -37.8085
27.4448 | -35.7396 | -37.8137
27.4509 | -35.7399 | -37.819
27.457 | -35.7403 | -37.8241
27.4631 | -35.7406 | -37.8296
27.4692 | -35.7409 | -37.8349
27.4753 | -35.7412 | -37.8395
27.4814 | -35.7415 | -37.8449
27.4875 | -35.7419 | -37.8502
27.4936 | -35.7423 | -37.8557
27.4997 | -35.7426 | -37.8611
27.5058 | -35.7429 | -37.8665
27.5119 | -35.7432 | -37.8718
27.518 | -35.7435 | -37.8769
27.5241 | -35.7438 | -37.882
27.5302 | -35.7442 | -37.8874
27.5363 | -35.7445 | -37.8928
27.5424 | -35.7448 | -37.8981
27.5485 | -35.7452 | -37.9036
27.5546 | -35.7455 | -37.9087
27.5607 | -35.7458 | -37.9141
27.5668 | -35.7462 | -37.9194
27.5729 | -35.7465 | -37.9246
27.5791 | -35.7469 | -37.9293
27.5852 | -35.7472 | -37.9345
27.5913 | -35.7475 | -37.9399
27.5974 | -35.7478 | -37.945
27.6035 | -35.7481 | -37.9503
27.6096 | -35.7484 | -37.9553
27.6157 | -35.7487 | -37.9605
27.6218 | -35.749 | -37.9656
27.6279 | -35.7494 | -37.9706
27.634 | -35.7497 | -37.9759
27.6401 | -35.75 | -37.9811
27.6462 | -35.7504 | -37.9865
27.6523 | -35.7507 | -37.9911
27.6584 | -35.7511 | -37.9962
27.6645 | -35.7513 | -38.0015
27.6706 | -35.7516 | -38.0067
27.6767 | -35.752 | -38.0121
27.6828 | -35.7523 | -38.0173
27.6889 | -35.7526 | -38.0224
27.695 | -35.7529 | -38.0275
27.7011 | -35.7533 | -38.0326
27.7072 | -35.7536 | -38.0382
27.7133 | -35.7539 | -38.0435
27.7194 | -35.7543 | -38.0484
27.7256 | -35.7546 | -38.0538
27.7317 | -35.7549 | -38.0589
27.7378 | -35.7552 | -38.064
27.7439 | -35.7555 | -38.0695
27.75 | -35.7558 | -38.0746
27.7561 | -35.7561 | -38.0794
27.7622 | -35.7564 | -38.0844
27.7683 | -35.7567 | -38.0894
27.7744 | -35.757 | -38.0947
27.7805 | -35.7573 | -38.0997
27.7866 | -35.7576 | -38.105
27.7927 | -35.7579 | -38.1101
27.7988 | -35.7582 | -38.1156
27.8049 | -35.7585 | -38.1208
27.811 | -35.7589 | -38.1258
27.8171 | -35.7592 | -38.1313
27.8232 | -35.7595 | -38.1361
27.8293 | -35.7598 | -38.1413
27.8354 | -35.7602 | -38.1467
27.8415 | -35.7605 | -38.1522
27.8476 | -35.7609 | -38.1575
27.8537 | -35.7612 | -38.1629
27.8598 | -35.7616 | -38.1678
27.866 | -35.7619 | -38.1731
27.8721 | -35.7622 | -38.1781
27.8782 | -35.7625 | -38.1832
27.8843 | -35.7629 | -38.1882
27.8904 | -35.7632 | -38.1934
27.8965 | -35.7634 | -38.1981
27.9026 | -35.7637 | -38.2035
27.9087 | -35.764 | -38.2088
27.9148 | -35.7644 | -38.214
27.9209 | -35.7647 | -38.2193
27.927 | -35.765 | -38.2243
27.9331 | -35.7654 | -38.2292
27.9392 | -35.7657 | -38.2346
27.9453 | -35.766 | -38.2399
27.9514 | -35.7663 | -38.2451
27.9575 | -35.7666 | -38.2503
27.9636 | -35.7669 | -38.2554
27.9697 | -35.7672 | -38.2603
27.9758 | -35.7676 | -38.2657
27.9819 | -35.7679 | -38.271
27.988 | -35.7682 | -38.2761
27.9941 | -35.7685 | -38.2814
28.0002 | -35.7688 | -38.2863
28.0063 | -35.7691 | -38.2913
28.0125 | -35.7694 | -38.2964
28.0186 | -35.7697 | -38.3017
28.0247 | -35.7701 | -38.307
28.0308 | -35.7704 | -38.3123
28.0369 | -35.7707 | -38.317
28.043 | -35.771 | -38.3221
28.0491 | -35.7713 | -38.3271
28.0552 | -35.7716 | -38.332
28.0613 | -35.7719 | -38.337
28.0674 | -35.7722 | -38.342
28.0735 | -35.7725 | -38.3467

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4356 | -42.1948
21.4908 | -35.5087 | -35.2075
16.4948 | -34.8812 | -27.7791
12.6602 | -34.5138 | -21.4835
9.7171 | -34.2981 | -16.5533
7.4582 | -34.1477 | -12.5646
5.7244 | -34.0377 | -9.3348
4.3936 | -33.9697 | -6.7268
3.3722 | -33.9485 | -4.6763
2.5883 | -33.9475 | -3.1172
1.9866 | -33.9478 | -1.8745
1.5248 | -33.951 | -0.87709
1.1703 | -33.9608 | -0.22787
1 | -33.9521 | 0.10039
0.89822 | -33.9531 | 0.36951
0.68941 | -33.9413 | 0.95263
0.52919 | -33.9365 | 1.3993
0.40612 | -33.9372 | 2.0091
0.31171 | -33.9384 | 2.4597
0.23928 | -33.9415 | 3.186
0.18358 | -33.9556 | 3.6036
0.14091 | -33.9507 | 4.491
0.10819 | -33.9862 | 5.6475
0.083017 | -33.9988 | 7.131
0.063705 | -34.0386 | 9.0098
0.048923 | -34.1177 | 11.5162
0.037479 | -34.2232 | 14.7206
0.028801 | -34.3984 | 18.7394
0.022078 | -34.717 | 24.0467
0.017023 | -35.1821 | 30.4106
0.013018 | -35.984 | 38.0632
0.010014 | -37.093 | 45.9682

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003743 | -10.1256
0.007485 | -34.3973
0.011228 | -38.1539
0.014971 | -40.4016
0.018713 | -43.8644
0.022456 | -45.8908
0.026199 | -46.9048
0.029942 | -49.6776
0.033684 | -51.1276
0.037427 | -52.7783
0.04117 | -54.058
0.044912 | -55.0363
0.048655 | -55.7649
0.052398 | -56.5101
0.05614 | -56.6079
0.059883 | -56.8875
0.063626 | -57.7011
0.067368 | -58.7731
0.071111 | -60.0056
0.074854 | -61.1086
0.078596 | -61.7632
0.082339 | -62.274
0.086082 | -62.4532
0.089825 | -62.165
0.093567 | -63.26
0.09731 | -64.3446
0.10105 | -64.807
0.10479 | -65.4088
0.10854 | -65.1798
0.11228 | -65.0409
0.11602 | -65.182
0.11977 | -65.2046
0.12351 | -65.6016
0.12725 | -66.5994
0.13099 | -67.7013
0.13474 | -67.6172
0.13848 | -69.0214
0.14222 | -68.4562
0.14597 | -68.5101
0.14971 | -69.2042
0.15345 | -69.4264
0.15719 | -69.0466
0.16094 | -68.6989
0.16468 | -68.5367
0.16842 | -68.7644
0.17216 | -70.1676
0.17591 | -69.9365
0.17965 | -71.4144
0.18339 | -70.8828
0.18713 | -70.6959
0.19088 | -71.1621
0.19462 | -71.6017
0.19836 | -71.836
0.20211 | -71.5241
0.20585 | -72.2934
0.20959 | -72.4476
0.21333 | -72.5502
0.21708 | -73.3405
0.22082 | -73.7457
0.22456 | -73.2327
0.2283 | -73.0688
0.23205 | -73.4221
0.23579 | -73.2166
0.23953 | -72.8364
0.24327 | -72.9708
0.24702 | -73.325
0.25076 | -72.5729
0.2545 | -73.1264
0.25825 | -73.882
0.26199 | -74.845
0.26573 | -75.3924
0.26947 | -76.0775
0.27322 | -75.1847
0.27696 | -74.8815
0.2807 | -75.3256
0.28444 | -75.2902
0.28819 | -74.9943
0.29193 | -74.602
0.29567 | -75.178
0.29941 | -75.9992
0.30316 | -76.3728
0.3069 | -76.9942
0.31064 | -76.9538
0.31439 | -76.407
0.31813 | -75.5557
0.32187 | -75.9502
0.32561 | -76.4472
0.32936 | -76.8779
0.3331 | -76.8292
0.33684 | -77.2587
0.34059 | -78.1545
0.34433 | -77.9368
0.34807 | -76.3978
0.35181 | -76.7205
0.35556 | -77.1246
0.3593 | -77.262
0.36304 | -77.5896
0.36678 | -78.0282
0.37053 | -77.1638
0.37427 | -77.3486
0.37801 | -77.2272
0.38175 | -77.2754
0.3855 | -77.8032
0.38924 | -77.8853
0.39298 | -78.633
0.39672 | -78.5725
0.40047 | -78.6786
0.40421 | -79.0079
0.40795 | -80.2547
0.4117 | -79.4531
0.41544 | -78.6174
0.41918 | -79.3761
0.42292 | -79.5996
0.42667 | -79.7995
0.43041 | -80.0685
0.43415 | -79.9695
0.43789 | -79.4753
0.44164 | -78.7941
0.44538 | -78.5867
0.44912 | -79.2499
0.45287 | -79.6044
0.45661 | -80.0004
0.46035 | -79.9333
0.46409 | -80.2713
0.46784 | -80.2034
0.47158 | -79.3709
0.47532 | -79.5442
0.47906 | -80.3812
0.48281 | -79.3191
0.48655 | -78.7384
0.49029 | -79.5277
0.49404 | -80.9883
0.49778 | -81.9345
0.50152 | -81.2535
0.50526 | -80.5126
0.50901 | -79.9545
0.51275 | -80.1517
0.51649 | -80.5567
0.52023 | -80.7907
0.52398 | -80.7251
0.52772 | -81.2769
0.53146 | -82.3201
0.53521 | -81.8221
0.53895 | -81.5804
0.54269 | -81.7652
0.54643 | -82.1896
0.55017 | -82.6178
0.55392 | -82.5115
0.55766 | -82.2289
0.5614 | -82.4017
0.56515 | -82.6204
0.56889 | -82.6515
0.57263 | -82.798
0.57637 | -83.1869
0.58012 | -83.3227
0.58386 | -83.0436
0.5876 | -82.7009
0.59135 | -82.2968
0.59509 | -82.3785
0.59883 | -82.84
0.60257 | -82.7315
0.60632 | -82.4953
0.61006 | -82.6214
0.6138 | -82.7218
0.61754 | -82.8447
0.62129 | -82.7606
0.62503 | -82.3567
0.62877 | -82.1571
0.63252 | -81.8979
0.63626 | -82.0497
0.64 | -82.4668
0.64374 | -82.7889
0.64748 | -82.8687
0.65123 | -82.7291
0.65497 | -82.5363
0.65871 | -82.2753
0.66246 | -81.9125
0.6662 | -81.7486
0.66994 | -82.1774
0.67368 | -82.6826
0.67743 | -82.6922
0.68117 | -82.7947
0.68491 | -83.4746
0.68866 | -84.189
0.6924 | -84.3621
0.69614 | -84.3331
0.69988 | -84.0785
0.70363 | -83.7906
0.70737 | -83.6734
0.71111 | -83.296
0.71485 | -83.143
0.7186 | -83.3191
0.72234 | -84.2173
0.72608 | -85.1359
0.72982 | -84.7478
0.73357 | -84.1269
0.73731 | -83.8612
0.74105 | -83.869
0.74479 | -84.2949
0.74854 | -84.7442
0.75228 | -85.0704
0.75602 | -85.3415
0.75977 | -85.6033
0.76351 | -86.0499
0.76725 | -85.83
0.77099 | -85.0293
0.77474 | -84.5143
0.77848 | -84.3339
0.78222 | -84.0233
0.78597 | -84.3296
0.78971 | -84.9528
0.79345 | -85.0382
0.79719 | -84.942
0.80094 | -84.853
0.80468 | -85.326
0.80842 | -85.6788
0.81216 | -85.7054
0.81591 | -85.8866
0.81965 | -85.6578
0.82339 | -84.9528
0.82713 | -84.5705
0.83088 | -84.7433
0.83462 | -84.9272
0.83836 | -85.1554
0.8421 | -85.4277
0.84585 | -85.6144
0.84959 | -86.0026
0.85333 | -86.5378
0.85708 | -86.819
0.86082 | -86.9287
0.86456 | -86.6492
0.8683 | -85.9599
0.87205 | -85.5008
0.87579 | -85.4474
0.87953 | -85.6751
0.88328 | -86.1165
0.88702 | -86.5289
0.89076 | -87.0157
0.8945 | -87.096
0.89825 | -86.6855
0.90199 | -86.7203
0.90573 | -86.9477
0.90947 | -86.9295
0.91322 | -86.7782
0.91696 | -86.5567
0.9207 | -85.8581
0.92444 | -85.487
0.92819 | -85.7223
0.93193 | -85.9514
0.93567 | -86.4814
0.93942 | -86.9374
0.94316 | -86.7322
0.9469 | -86.2339
0.95064 | -86.3615
0.95439 | -86.9996
0.95813 | -86.8413
0.96187 | -86.3889
0.96561 | -86.2945
0.96936 | -86.1973
0.9731 | -86.4156
0.97684 | -86.8913
0.98059 | -87.4235
0.98433 | -87.5439
0.98807 | -87.3543
0.99181 | -87.587
0.99556 | -87.6254
0.9993 | -86.9484
1.003 | -86.0055
1.0068 | -85.6449
1.0105 | -86.3056
1.0143 | -86.9601
1.018 | -87.1824
1.0218 | -87.2761
1.0255 | -87.0599
1.0292 | -87.1175
1.033 | -87.3894
1.0367 | -87.6626
1.0405 | -87.9529
1.0442 | -87.7523
1.048 | -86.9023
1.0517 | -86.2234
1.0554 | -86.6156
1.0592 | -87.4507
1.0629 | -88.1375
1.0667 | -88.3451
1.0704 | -87.9868
1.0742 | -87.4664
1.0779 | -87.3078
1.0816 | -87.3315
1.0854 | -87.5106
1.0891 | -87.6507
1.0929 | -87.7098
1.0966 | -87.8238
1.1004 | -87.9814
1.1041 | -88.0208
1.1078 | -87.933
1.1116 | -87.756
1.1153 | -87.5924
1.1191 | -87.5727
1.1228 | -87.6082
1.1265 | -87.5437
1.1303 | -87.5176
1.134 | -87.6663
1.1378 | -87.8441
1.1415 | -87.8529
1.1453 | -87.7035
1.149 | -87.5741
1.1527 | -87.6041
1.1565 | -87.8004
1.1602 | -88.0105
1.164 | -88.0421
1.1677 | -88.0303
1.1715 | -88.176
1.1752 | -88.3574
1.1789 | -88.4136
1.1827 | -88.3878
1.1864 | -88.371
1.1902 | -88.3746
1.1939 | -88.3986
1.1977 | -88.5323
1.2014 | -88.853
1.2051 | -89.2236
1.2089 | -89.3905
1.2126 | -89.4027
1.2164 | -89.3826
1.2201 | -89.3091
1.2239 | -89.2791
1.2276 | -89.408
1.2313 | -89.5935
1.2351 | -89.6068
1.2388 | -89.3144
1.2426 | -88.8425
1.2463 | -88.5115
1.2501 | -88.5437
1.2538 | -88.7532
1.2575 | -88.7894
1.2613 | -88.7363
1.265 | -88.8956
1.2688 | -89.1547
1.2725 | -89.1405
1.2763 | -88.8955
1.28 | -88.7913
1.2837 | -88.9038
1.2875 | -89.0622
1.2912 | -89.2238
1.295 | -89.4077
1.2987 | -89.4325
1.3025 | -89.2033
1.3062 | -88.9518
1.3099 | -88.9848
1.3137 | -89.3318
1.3174 | -89.5811
1.3212 | -89.3924
1.3249 | -89.1379
1.3287 | -89.2475
1.3324 | -89.6552
1.3361 | -90.0173
1.3399 | -90.1529
1.3436 | -90.1497
1.3474 | -90.2241
1.3511 | -90.4275
1.3549 | -90.3848
1.3586 | -89.9271
1.3623 | -89.619
1.3661 | -89.7496
1.3698 | -89.9474
1.3736 | -89.9305
1.3773 | -89.9744
1.3811 | -90.3019
1.3848 | -90.7216
1.3885 | -90.9143
1.3923 | -90.9332
1.396 | -90.8966
1.3998 | -90.765
1.4035 | -90.5959
1.4073 | -90.378
1.411 | -90.0651
1.4147 | -89.8259
1.4185 | -89.8073
1.4222 | -89.9773
1.426 | -90.2034
1.4297 | -90.3334
1.4334 | -90.4147
1.4372 | -90.5696
1.4409 | -90.6311
1.4447 | -90.4327
1.4484 | -90.2236
1.4522 | -90.2119
1.4559 | -90.3072
1.4596 | -90.4493
1.4634 | -90.7186
1.4671 | -91.1135
1.4709 | -91.3889
1.4746 | -91.2405
1.4784 | -90.8013
1.4821 | -90.4741
1.4858 | -90.4596
1.4896 | -90.6023
1.4933 | -90.5389
1.4971 | -90.3075
1.5008 | -90.1064
1.5046 | -89.9149
1.5083 | -89.8173
1.512 | -89.9746
1.5158 | -90.2989
1.5195 | -90.4766
1.5233 | -90.3322
1.527 | -90.1693
1.5308 | -90.4012
1.5345 | -91.0859
1.5382 | -91.799
1.542 | -92.0452
1.5457 | -91.841
1.5495 | -91.5142
1.5532 | -91.2411
1.557 | -91.0514
1.5607 | -91.0819
1.5644 | -91.3914
1.5682 | -91.5358
1.5719 | -91.155
1.5757 | -90.5773
1.5794 | -90.1364
1.5832 | -89.7984
1.5869 | -89.5734
1.5906 | -89.711
1.5944 | -90.2277
1.5981 | -90.7046
1.6019 | -90.8228
1.6056 | -90.7924
1.6094 | -90.9626
1.6131 | -91.3436
1.6168 | -91.6058
1.6206 | -91.7582
1.6243 | -91.7036
1.6281 | -91.4546
1.6318 | -91.1201
1.6356 | -90.8704
1.6393 | -90.8561
1.643 | -91.1437
1.6468 | -91.6027
1.6505 | -91.9226
1.6543 | -91.8911
1.658 | -91.6029
1.6618 | -91.2758
1.6655 | -91.076
1.6692 | -91.0355
1.673 | -91.1069
1.6767 | -91.2398
1.6805 | -91.4038
1.6842 | -91.56
1.688 | -91.6807
1.6917 | -91.7912
1.6954 | -91.8667
1.6992 | -91.8727
1.7029 | -91.8456
1.7067 | -91.8248
1.7104 | -91.7919
1.7142 | -91.7523
1.7179 | -91.7224
1.7216 | -91.7075
1.7254 | -91.7199
1.7291 | -91.7767
1.7329 | -91.8402
1.7366 | -91.8532
1.7404 | -91.8085
1.7441 | -91.7226
1.7478 | -91.6316
1.7516 | -91.6202
1.7553 | -91.7567
1.7591 | -91.9799
1.7628 | -92.1029
1.7666 | -92.0418
1.7703 | -91.8627
1.774 | -91.7061
1.7778 | -91.6712
1.7815 | -91.79
1.7853 | -92.0238
1.789 | -92.2574
1.7927 | -92.3716
1.7965 | -92.3437
1.8002 | -92.2399
1.804 | -92.1732
1.8077 | -92.1293
1.8115 | -92.0801
1.8152 | -92.0528
1.8189 | -92.1035
1.8227 | -92.2832
1.8264 | -92.5872
1.8302 | -92.9086
1.8339 | -93.0728
1.8377 | -93.0168
1.8414 | -92.8366
1.8451 | -92.5873
1.8489 | -92.3004
1.8526 | -92.0367
1.8564 | -91.8046
1.8601 | -91.5994
1.8639 | -91.446
1.8676 | -91.3989
1.8713 | -91.4661
1.8751 | -91.6016
1.8788 | -91.7403
1.8826 | -91.8818
1.8863 | -92.0979
1.8901 | -92.4109
1.8938 | -92.6796
1.8975 | -92.7736
1.9013 | -92.7583
1.905 | -92.7735
1.9088 | -92.8247
1.9125 | -92.8061
1.9163 | -92.6323
1.92 | -92.3805
1.9237 | -92.2023
1.9275 | -92.1743
1.9312 | -92.2855
1.935 | -92.4956
1.9387 | -92.7474
1.9425 | -92.9526
1.9462 | -93.0449
1.9499 | -93.0307
1.9537 | -92.9701
1.9574 | -92.9225
1.9612 | -92.8851
1.9649 | -92.7916
1.9687 | -92.6451
1.9724 | -92.4719
1.9761 | -92.2977
1.9799 | -92.1969
1.9836 | -92.2583
1.9874 | -92.4603
1.9911 | -92.6832
1.9949 | -92.8522
1.9986 | -92.9712
2.0023 | -93.0649
2.0061 | -93.1134
2.0098 | -93.0766
2.0136 | -92.9674
2.0173 | -92.8494
2.0211 | -92.7518
2.0248 | -92.6772
2.0285 | -92.6733
2.0323 | -92.7914
2.036 | -92.9834
2.0398 | -93.1574
2.0435 | -93.2698
2.0473 | -93.3802
2.051 | -93.5513
2.0547 | -93.7624
2.0585 | -93.9381
2.0622 | -94.0296
2.066 | -94.062
2.0697 | -94.0353
2.0734 | -93.9394
2.0772 | -93.7865
2.0809 | -93.6549
2.0847 | -93.6281
2.0884 | -93.7284
2.0922 | -93.8939
2.0959 | -94.0141
2.0996 | -93.9671
2.1034 | -93.7257
2.1071 | -93.4118
2.1109 | -93.205
2.1146 | -93.1445
2.1184 | -93.1633
2.1221 | -93.2021
2.1258 | -93.2645
2.1296 | -93.3494
2.1333 | -93.4457
2.1371 | -93.5235
2.1408 | -93.5704
2.1446 | -93.6295
2.1483 | -93.7884
2.152 | -94.0603
2.1558 | -94.3282
2.1595 | -94.366
2.1633 | -94.3996
2.167 | -94.3729
2.1708 | -94.2922
2.1745 | -94.1798
2.1782 | -94.0458
2.182 | -93.9278
2.1857 | -93.8804
2.1895 | -93.9156
2.1932 | -93.9958
2.197 | -94.0552
2.2007 | -94.047
2.2044 | -93.9864
2.2082 | -93.9159
2.2119 | -93.8705
2.2157 | -93.8562
2.2194 | -93.8836
2.2232 | -93.9486
2.2269 | -94.0407
2.2306 | -94.1415
2.2344 | -94.2361
2.2381 | -94.3067
2.2419 | -94.3386
2.2456 | -94.2893
2.2494 | -94.1335
2.2531 | -93.902
2.2568 | -93.6638
2.2606 | -93.5059
2.2643 | -93.479
2.2681 | -93.5743
2.2718 | -93.7377
2.2756 | -93.8914
2.2793 | -93.9802
2.283 | -94.0132
2.2868 | -94.0153
2.2905 | -93.9973
2.2943 | -93.938
2.298 | -93.8546
2.3018 | -93.7533
2.3055 | -93.6429
2.3092 | -93.5661
2.313 | -93.5604
2.3167 | -93.652
2.3205 | -93.8327
2.3242 | -94.0341
2.328 | -94.2087
2.3317 | -94.3459
2.3354 | -94.4597
2.3392 | -94.512
2.3429 | -94.4658
2.3467 | -94.2956
2.3504 | -94.065
2.3542 | -93.8757
2.3579 | -93.7906
2.3616 | -93.822
2.3654 | -93.9585
2.3691 | -94.158
2.3729 | -94.3446
2.3766 | -94.4566
2.3804 | -94.4607
2.3841 | -94.3912
2.3878 | -94.2869
2.3916 | -94.1664
2.3953 | -94.0775
2.3991 | -94.0551
2.4028 | -94.0982
2.4066 | -94.1754
2.4103 | -94.2324
2.414 | -94.2458
2.4178 | -94.2396
2.4215 | -94.2531
2.4253 | -94.2888
2.429 | -94.3154
2.4327 | -94.3081
2.4365 | -94.2723
2.4402 | -94.2458
2.444 | -94.2452
2.4477 | -94.2433
2.4515 | -94.2052
2.4552 | -94.1238
2.4589 | -94.0101
2.4627 | -93.9168
2.4664 | -93.8856
2.4702 | -93.9391
2.4739 | -94.0404
2.4777 | -94.1515
2.4814 | -94.2414
2.4851 | -94.3273
2.4889 | -94.4526
2.4926 | -94.6558
2.4964 | -94.912
2.5001 | -95.1645
2.5039 | -95.3347
2.5076 | -95.4023
2.5113 | -95.3964
2.5151 | -95.353
2.5188 | -95.2948
2.5226 | -95.2592
2.5263 | -95.2428
2.5301 | -95.2308
2.5338 | -95.2051
2.5375 | -95.1744
2.5413 | -95.1731
2.545 | -95.2429
2.5488 | -95.3775
2.5525 | -95.5366
2.5563 | -95.6246
2.56 | -95.5696
2.5637 | -95.3806
2.5675 | -95.1318
2.5712 | -94.9081
2.575 | -94.7843
2.5787 | -94.8025
2.5825 | -94.9165
2.5862 | -95.0453
2.5899 | -95.1288
2.5937 | -95.143
2.5974 | -95.1263
2.6012 | -95.0863
2.6049 | -95.0065
2.6087 | -94.8954
2.6124 | -94.7924
2.6161 | -94.7786
2.6199 | -94.902
2.6236 | -95.1538
2.6274 | -95.4384
2.6311 | -95.6422
2.6349 | -95.7389
2.6386 | -95.7596
2.6423 | -95.7424
2.6461 | -95.6852
2.6498 | -95.5864
2.6536 | -95.4795
2.6573 | -95.3761
2.6611 | -95.275
2.6648 | -95.1502
2.6685 | -95.0003
2.6723 | -94.8866
2.676 | -94.8536
2.6798 | -94.9038
2.6835 | -94.986
2.6873 | -95.0267
2.691 | -95.0124
2.6947 | -94.9943
2.6985 | -95.0549
2.7022 | -95.1115
2.706 | -95.1612
2.7097 | -95.183
2.7134 | -95.1822
2.7172 | -95.1523
2.7209 | -95.0941
2.7247 | -95.0178
2.7284 | -94.9421
2.7322 | -94.8998
2.7359 | -94.8917
2.7396 | -94.914
2.7434 | -94.9555
2.7471 | -95.0152
2.7509 | -95.0985
2.7546 | -95.219
2.7584 | -95.3667
2.7621 | -95.5182
2.7658 | -95.6442
2.7696 | -95.7341
2.7733 | -95.7959
2.7771 | -95.8477
2.7808 | -95.9005
2.7846 | -95.9507
2.7883 | -95.9876
2.792 | -95.9975
2.7958 | -95.9832
2.7995 | -95.9653
2.8033 | -95.9489
2.807 | -95.9339
2.8108 | -95.912
2.8145 | -95.8603
2.8182 | -95.7507
2.822 | -95.5697
2.8257 | -95.3405
2.8295 | -95.1211
2.8332 | -94.978
2.837 | -94.9384
2.8407 | -95.0096
2.8444 | -95.1446
2.8482 | -95.2885
2.8519 | -95.4021
2.8557 | -95.4569
2.8594 | -95.4631
2.8632 | -95.45
2.8669 | -95.4416
2.8706 | -95.4479
2.8744 | -95.4667
2.8781 | -95.4833
2.8819 | -95.505
2.8856 | -95.5453
2.8894 | -95.6099
2.8931 | -95.7099
2.8968 | -95.8223
2.9006 | -95.923
2.9043 | -95.9867
2.9081 | -96.0074
2.9118 | -95.9769
2.9156 | -95.9045
2.9193 | -95.831
2.923 | -95.793
2.9268 | -95.8157
2.9305 | -95.9048
2.9343 | -96.0302
2.938 | -96.1483
2.9418 | -96.2177
2.9455 | -96.2419
2.9492 | -96.2348
2.953 | -96.198
2.9567 | -96.119
2.9605 | -96.0297
2.9642 | -95.9699
2.968 | -95.9653
2.9717 | -96.0139
2.9754 | -96.0993
2.9792 | -96.1636
2.9829 | -96.1893
2.9867 | -96.1589
2.9904 | -96.0888
2.9942 | -96.0015
2.9979 | -95.9191
3.0016 | -95.8573
3.0054 | -95.8073
3.0091 | -95.7583
3.0129 | -95.7077
3.0166 | -95.6792
3.0204 | -95.6951
3.0241 | -95.7237
3.0278 | -95.7527
3.0316 | -95.7753
3.0353 | -95.8044
3.0391 | -95.8714
3.0428 | -95.9788
3.0465 | -96.0804
3.0503 | -96.1382
3.054 | -96.132
3.0578 | -96.0457
3.0615 | -95.906
3.0653 | -95.7428
3.069 | -95.5975
3.0727 | -95.5374
3.0765 | -95.5828
3.0802 | -95.7259
3.084 | -95.9401
3.0877 | -96.1505
3.0915 | -96.3232
3.0952 | -96.4478
3.0989 | -96.5072
3.1027 | -96.5275
3.1064 | -96.5626
3.1102 | -96.624
3.1139 | -96.7136
3.1177 | -96.8139
3.1214 | -96.9099
3.1251 | -96.9696
3.1289 | -96.9591
3.1326 | -96.892
3.1364 | -96.789
3.1401 | -96.6654
3.1439 | -96.5656
3.1476 | -96.5026
3.1513 | -96.4796
3.1551 | -96.4795
3.1588 | -96.4853
3.1626 | -96.4904
3.1663 | -96.4861
3.1701 | -96.4627
3.1738 | -96.4471
3.1775 | -96.4548
3.1813 | -96.4908
3.185 | -96.5798
3.1888 | -96.7205
3.1925 | -96.8757
3.1963 | -97.0157
3.2 | -97.0993
3.2037 | -97.1122
3.2075 | -97.0677
3.2112 | -96.972
3.215 | -96.8633
3.2187 | -96.7775
3.2225 | -96.7236
3.2262 | -96.7054
3.2299 | -96.716
3.2337 | -96.7644
3.2374 | -96.8241
3.2412 | -96.8674
3.2449 | -96.9123
3.2487 | -96.9498
3.2524 | -96.9768
3.2561 | -97.0002
3.2599 | -97.0255
3.2636 | -97.0608
3.2674 | -97.0941
3.2711 | -97.1257
3.2749 | -97.1742
3.2786 | -97.2176
3.2823 | -97.2382
3.2861 | -97.2263
3.2898 | -97.2093
3.2936 | -97.1858
3.2973 | -97.1492
3.3011 | -97.1096
3.3048 | -97.0539
3.3085 | -96.9814
3.3123 | -96.902
3.316 | -96.8188
3.3198 | -96.7523
3.3235 | -96.7204
3.3273 | -96.7198
3.331 | -96.7357
3.3347 | -96.771
3.3385 | -96.8126
3.3422 | -96.8516
3.346 | -96.8926
3.3497 | -96.9412
3.3535 | -97.0033
3.3572 | -97.071
3.3609 | -97.1371
3.3647 | -97.1889
3.3684 | -97.2216
3.3722 | -97.221
3.3759 | -97.1744
3.3796 | -97.0886
3.3834 | -96.979
3.3871 | -96.8691
3.3909 | -96.7927
3.3946 | -96.775
3.3984 | -96.8089
3.4021 | -96.871
3.4058 | -96.9471
3.4096 | -97.0143
3.4133 | -97.0587
3.4171 | -97.0953
3.4208 | -97.1262
3.4246 | -97.1485
3.4283 | -97.1636
3.432 | -97.1713
3.4358 | -97.1727
3.4395 | -97.169
3.4433 | -97.1555
3.447 | -97.1368
3.4508 | -97.106
3.4545 | -97.0767
3.4582 | -97.0733
3.462 | -97.0953
3.4657 | -97.1378
3.4695 | -97.1949
3.4732 | -97.2521
3.477 | -97.3077
3.4807 | -97.3604
3.4844 | -97.4187
3.4882 | -97.4706
3.4919 | -97.5139
3.4957 | -97.5445
3.4994 | -97.5695
3.5032 | -97.5901
3.5069 | -97.6067
3.5106 | -97.6113
3.5144 | -97.5895
3.5181 | -97.5452
3.5219 | -97.4874
3.5256 | -97.4246
3.5294 | -97.3568
3.5331 | -97.291
3.5368 | -97.2528
3.5406 | -97.2482
3.5443 | -97.2779
3.5481 | -97.3353
3.5518 | -97.4118
3.5556 | -97.4969
3.5593 | -97.5716
3.563 | -97.646
3.5668 | -97.7188
3.5705 | -97.7858
3.5743 | -97.8473
3.578 | -97.9054
3.5818 | -97.9437
3.5855 | -97.965
3.5892 | -97.9491
3.593 | -97.8728
3.5967 | -97.7487
3.6005 | -97.611
3.6042 | -97.4953
3.608 | -97.4219
3.6117 | -97.4052
3.6154 | -97.4365
3.6192 | -97.4864
3.6229 | -97.5273
3.6267 | -97.5442
3.6304 | -97.5349
3.6342 | -97.5248
3.6379 | -97.5364
3.6416 | -97.5693
3.6454 | -97.6279
3.6491 | -97.6922
3.6529 | -97.7426
3.6566 | -97.7695
3.6604 | -97.7659
3.6641 | -97.7397
3.6678 | -97.6822
3.6716 | -97.614
3.6753 | -97.5497
3.6791 | -97.5005
3.6828 | -97.4704
3.6865 | -97.4493
3.6903 | -97.4406
3.694 | -97.4558
3.6978 | -97.4996
3.7015 | -97.5814
3.7053 | -97.6919
3.709 | -97.809
3.7127 | -97.9134
3.7165 | -97.9812
3.7202 | -98.0101
3.724 | -98.0138
3.7277 | -97.9909
3.7315 | -97.9424
3.7352 | -97.8855
3.7389 | -97.8265
3.7427 | -97.7679
3.7464 | -97.711
3.7502 | -97.68
3.7539 | -97.6694
3.7577 | -97.6653
3.7614 | -97.688
3.7651 | -97.7188
3.7689 | -97.7531
3.7726 | -97.8007
3.7764 | -97.8634
3.7801 | -97.9269
3.7839 | -97.9957
3.7876 | -98.0659
3.7913 | -98.1328
3.7951 | -98.1802
3.7988 | -98.2142
3.8026 | -98.232
3.8063 | -98.219
3.8101 | -98.1777
3.8138 | -98.1263
3.8175 | -98.0717
3.8213 | -98.009
3.825 | -97.9539
3.8288 | -97.91
3.8325 | -97.8788
3.8363 | -97.8524
3.84 | -97.8234
3.8437 | -97.7909
3.8475 | -97.7521
3.8512 | -97.7155
3.855 | -97.688
3.8587 | -97.6692
3.8625 | -97.654
3.8662 | -97.6456
3.8699 | -97.6512
3.8737 | -97.6672
3.8774 | -97.7011
3.8812 | -97.7485
3.8849 | -97.8165
3.8887 | -97.91
3.8924 | -98.0155
3.8961 | -98.1119
3.8999 | -98.1999
3.9036 | -98.2709
3.9074 | -98.3203
3.9111 | -98.3479
3.9149 | -98.3543
3.9186 | -98.3399
3.9223 | -98.3201
3.9261 | -98.2901
3.9298 | -98.2447
3.9336 | -98.1906
3.9373 | -98.133
3.9411 | -98.0721
3.9448 | -98.025
3.9485 | -97.9907
3.9523 | -97.9741
3.956 | -97.9873
3.9598 | -98.0476
3.9635 | -98.1414
3.9673 | -98.2552
3.971 | -98.3794
3.9747 | -98.508
3.9785 | -98.6189
3.9822 | -98.7079
3.986 | -98.7758
3.9897 | -98.8175
3.9935 | -98.8319
3.9972 | -98.8068
4.0009 | -98.7567
4.0047 | -98.6704
4.0084 | -98.551
4.0122 | -98.4256
4.0159 | -98.2992
4.0196 | -98.1913
4.0234 | -98.1049
4.0271 | -98.0427
4.0309 | -98.0078
4.0346 | -97.9926
4.0384 | -97.9878
4.0421 | -98.0005
4.0458 | -98.031
4.0496 | -98.0683
4.0533 | -98.1072
4.0571 | -98.1384
4.0608 | -98.1416
4.0646 | -98.13
4.0683 | -98.1247
4.072 | -98.1199
4.0758 | -98.1093
4.0795 | -98.1134
4.0833 | -98.127
4.087 | -98.1403
4.0908 | -98.1564
4.0945 | -98.1716
4.0982 | -98.1875
4.102 | -98.2187
4.1057 | -98.2591
4.1095 | -98.3034
4.1132 | -98.3615
4.117 | -98.4243
4.1207 | -98.4678
4.1244 | -98.495
4.1282 | -98.4835
4.1319 | -98.4367
4.1357 | -98.3721
4.1394 | -98.303
4.1432 | -98.2293
4.1469 | -98.1641
4.1506 | -98.1282
4.1544 | -98.1283
4.1581 | -98.1674
4.1619 | -98.2416
4.1656 | -98.3497
4.1694 | -98.4887
4.1731 | -98.6419
4.1768 | -98.7915
4.1806 | -98.9357
4.1843 | -99.0728
4.1881 | -99.1864
4.1918 | -99.273
4.1956 | -99.3363
4.1993 | -99.3746
4.203 | -99.3888
4.2068 | -99.3903
4.2105 | -99.3853
4.2143 | -99.3691
4.218 | -99.3437
4.2218 | -99.3194
4.2255 | -99.2978
4.2292 | -99.2763
4.233 | -99.2592
4.2367 | -99.2455
4.2405 | -99.2303
4.2442 | -99.1978
4.248 | -99.1486
4.2517 | -99.0894
4.2554 | -99.0272
4.2592 | -98.9582
4.2629 | -98.888
4.2667 | -98.8264
4.2704 | -98.7841
4.2742 | -98.7733
4.2779 | -98.8038
4.2816 | -98.8757
4.2854 | -98.9685
4.2891 | -99.066
4.2929 | -99.1494
4.2966 | -99.1983
4.3004 | -99.2183
4.3041 | -99.2119
4.3078 | -99.1937
4.3116 | -99.1602
4.3153 | -99.1257
4.3191 | -99.1219
4.3228 | -99.1115
4.3266 | -99.1141
4.3303 | -99.1311
4.334 | -99.1611
4.3378 | -99.202
4.3415 | -99.2492
4.3453 | -99.3002
4.349 | -99.3452
4.3527 | -99.3924
4.3565 | -99.4373
4.3602 | -99.4758
4.364 | -99.5106
4.3677 | -99.5314
4.3715 | -99.5359
4.3752 | -99.5196
4.3789 | -99.4908
4.3827 | -99.4369
4.3864 | -99.351
4.3902 | -99.2374
4.3939 | -99.1113
4.3977 | -98.9827
4.4014 | -98.8581
4.4051 | -98.7521
4.4089 | -98.672
4.4126 | -98.6168
4.4164 | -98.597
4.4201 | -98.6079
4.4239 | -98.6495
4.4276 | -98.7159
4.4313 | -98.7955
4.4351 | -98.8838
4.4388 | -98.963
4.4426 | -99.0328
4.4463 | -99.0915
4.4501 | -99.1309
4.4538 | -99.1489
4.4575 | -99.1576
4.4613 | -99.1583
4.465 | -99.1463
4.4688 | -99.1042
4.4725 | -99.0511
4.4763 | -98.9949
4.48 | -98.943
4.4837 | -98.9059
4.4875 | -98.8849
4.4912 | -98.9009
4.495 | -98.9422
4.4987 | -99.003
4.5025 | -99.0803
4.5062 | -99.1575
4.5099 | -99.2242
4.5137 | -99.2636
4.5174 | -99.2662
4.5212 | -99.2433
4.5249 | -99.2133
4.5287 | -99.1797
4.5324 | -99.1516
4.5361 | -99.1505
4.5399 | -99.1898
4.5436 | -99.2619
4.5474 | -99.3613
4.5511 | -99.4786
4.5549 | -99.5984
4.5586 | -99.6874
4.5623 | -99.7339
4.5661 | -99.7331
4.5698 | -99.6827
4.5736 | -99.5943
4.5773 | -99.4832
4.5811 | -99.3778
4.5848 | -99.2998
4.5885 | -99.2506
4.5923 | -99.2398
4.596 | -99.2643
4.5998 | -99.3225
4.6035 | -99.4017
4.6073 | -99.4864
4.611 | -99.5824
4.6147 | -99.6701
4.6185 | -99.7427
4.6222 | -99.8049
4.626 | -99.8567
4.6297 | -99.8988
4.6334 | -99.9355
4.6372 | -99.9793
4.6409 | -100.0298
4.6447 | -100.0763
4.6484 | -100.1254
4.6522 | -100.1654
4.6559 | -100.1812
4.6596 | -100.1924
4.6634 | -100.1869
4.6671 | -100.1714
4.6709 | -100.1529
4.6746 | -100.1311
4.6784 | -100.1035
4.6821 | -100.0594
4.6858 | -99.9944
4.6896 | -99.9187
4.6933 | -99.8319
4.6971 | -99.7494
4.7008 | -99.6887
4.7046 | -99.6485
4.7083 | -99.6345
4.712 | -99.647
4.7158 | -99.681
4.7195 | -99.7258
4.7233 | -99.7848
4.727 | -99.8506
4.7308 | -99.9097
4.7345 | -99.9675
4.7382 | -100.025
4.742 | -100.0748
4.7457 | -100.1217
4.7495 | -100.1429
4.7532 | -100.1433
4.757 | -100.1319
4.7607 | -100.0976
4.7644 | -100.0464
4.7682 | -99.9862
4.7719 | -99.9166
4.7757 | -99.8514
4.7794 | -99.7967
4.7832 | -99.7536
4.7869 | -99.721
4.7906 | -99.6982
4.7944 | -99.6767
4.7981 | -99.6709
4.8019 | -99.686
4.8056 | -99.7094
4.8094 | -99.7293
4.8131 | -99.742
4.8168 | -99.7439
4.8206 | -99.737
4.8243 | -99.7167
4.8281 | -99.6848
4.8318 | -99.6503
4.8356 | -99.6209
4.8393 | -99.5972
4.843 | -99.5942
4.8468 | -99.6063
4.8505 | -99.6449
4.8543 | -99.7011
4.858 | -99.8084
4.8618 | -99.8926
4.8655 | -99.993
4.8692 | -100.1009
4.873 | -100.1958
4.8767 | -100.2747
4.8805 | -100.3327
4.8842 | -100.3717
4.888 | -100.3839
4.8917 | -100.3648
4.8954 | -100.3264
4.8992 | -100.263
4.9029 | -100.1887
4.9067 | -100.1254
4.9104 | -100.0676
4.9142 | -100.0247
4.9179 | -99.9993
4.9216 | -99.9859
4.9254 | -99.9894
4.9291 | -100.0091
4.9329 | -100.0297
4.9366 | -100.0407
4.9404 | -100.0529
4.9441 | -100.0526
4.9478 | -100.0453
4.9516 | -100.0353
4.9553 | -100.0153
4.9591 | -99.9858
4.9628 | -99.95
4.9665 | -99.9038
4.9703 | -99.8592
4.974 | -99.8173
4.9778 | -99.7836
4.9815 | -99.7482
4.9853 | -99.7133
4.989 | -99.6905
4.9927 | -99.6736
4.9965 | -99.6661
5.0002 | -99.6719
5.004 | -99.6785
5.0077 | -99.6941
5.0115 | -99.7199
5.0152 | -99.7542
5.0189 | -99.7925
5.0227 | -99.8376
5.0264 | -99.8852
5.0302 | -99.9287
5.0339 | -99.9743
5.0377 | -100.0159
5.0414 | -100.0528
5.0451 | -100.0953
5.0489 | -100.1363
5.0526 | -100.1774
5.0564 | -100.2249
5.0601 | -100.2779
5.0639 | -100.3238
5.0676 | -100.3534
5.0713 | -100.3728
5.0751 | -100.3826
5.0788 | -100.3845
5.0826 | -100.3795
5.0863 | -100.3694
5.0901 | -100.3555
5.0938 | -100.3447
5.0975 | -100.3331
5.1013 | -100.318
5.105 | -100.3042
5.1088 | -100.2982
5.1125 | -100.2906
5.1163 | -100.2937
5.12 | -100.3025
5.1237 | -100.3161
5.1275 | -100.3354
5.1312 | -100.3538
5.135 | -100.3594
5.1387 | -100.3551
5.1425 | -100.3394
5.1462 | -100.3154
5.1499 | -100.2727
5.1537 | -100.2176
5.1574 | -100.1597
5.1612 | -100.1111
5.1649 | -100.0799
5.1687 | -100.06
5.1724 | -100.046
5.1761 | -100.0532
5.1799 | -100.0727
5.1836 | -100.1038
5.1874 | -100.1427
5.1911 | -100.1876
5.1949 | -100.2241
5.1986 | -100.2567
5.2023 | -100.2861
5.2061 | -100.3085
5.2098 | -100.3273
5.2136 | -100.3436
5.2173 | -100.3567
5.2211 | -100.3574
5.2248 | -100.3508
5.2285 | -100.3443
5.2323 | -100.3352
5.236 | -100.3192
5.2398 | -100.2942
5.2435 | -100.2643
5.2473 | -100.2421
5.251 | -100.2333
5.2547 | -100.2255
5.2585 | -100.237
5.2622 | -100.2657
5.266 | -100.3055
5.2697 | -100.3573
5.2735 | -100.419
5.2772 | -100.4781
5.2809 | -100.5293
5.2847 | -100.5717
5.2884 | -100.5999
5.2922 | -100.6108
5.2959 | -100.6029
5.2996 | -100.599
5.3034 | -100.5903
5.3071 | -100.5791
5.3109 | -100.5758
5.3146 | -100.5911
5.3184 | -100.6185
5.3221 | -100.6521
5.3258 | -100.6926
5.3296 | -100.7357
5.3333 | -100.7712
5.3371 | -100.7857
5.3408 | -100.7848
5.3446 | -100.7769
5.3483 | -100.7565
5.352 | -100.7231
5.3558 | -100.6825
5.3595 | -100.6477
5.3633 | -100.6216
5.367 | -100.6103
5.3708 | -100.6148
5.3745 | -100.6333
5.3782 | -100.6615
5.382 | -100.7044
5.3857 | -100.7597
5.3895 | -100.8208
5.3932 | -100.8841
5.397 | -100.925
5.4007 | -100.9765
5.4044 | -101.0176
5.4082 | -101.0417
5.4119 | -101.0465
5.4157 | -101.0416
5.4194 | -101.0304
5.4232 | -100.9995
5.4269 | -100.9649
5.4306 | -100.9312
5.4344 | -100.9078
5.4381 | -100.8956
5.4419 | -100.8936
5.4456 | -100.91
5.4494 | -100.9374
5.4531 | -100.9785
5.4568 | -101.0316
5.4606 | -101.091
5.4643 | -101.1526
5.4681 | -101.2142
5.4718 | -101.2652
5.4756 | -101.2975
5.4793 | -101.3189
5.483 | -101.3251
5.4868 | -101.2946
5.4905 | -101.2427
5.4943 | -101.1658
5.498 | -101.0666
5.5018 | -100.9626
5.5055 | -100.8588
5.5092 | -100.7668
5.513 | -100.6944
5.5167 | -100.6539
5.5205 | -100.6432
5.5242 | -100.6581
5.528 | -100.6915
5.5317 | -100.7474
5.5354 | -100.8192
5.5392 | -100.8909
5.5429 | -100.965
5.5467 | -101.0404
5.5504 | -101.1005
5.5542 | -101.1354
5.5579 | -101.1538
5.5616 | -101.1588
5.5654 | -101.1311
5.5691 | -101.0898
5.5729 | -101.0422
5.5766 | -100.9839
5.5804 | -100.9307
5.5841 | -100.8815
5.5878 | -100.8271
5.5916 | -100.7787
5.5953 | -100.746
5.5991 | -100.7266
5.6028 | -100.7238
5.6066 | -100.7376
5.6103 | -100.7631
5.614 | -100.8115
5.6178 | -100.8625
5.6215 | -100.9119
5.6253 | -100.9667
5.629 | -101.0111
5.6327 | -101.0383
5.6365 | -101.0608
5.6402 | -101.0747
5.644 | -101.0755
5.6477 | -101.063
5.6515 | -101.0531
5.6552 | -101.0374
5.6589 | -101.032
5.6627 | -101.0455
5.6664 | -101.0575
5.6702 | -101.0736
5.6739 | -101.0844
5.6777 | -101.0862
5.6814 | -101.0768
5.6851 | -101.0517
5.6889 | -101.0167
5.6926 | -100.9795
5.6964 | -100.9461
5.7001 | -100.9025
5.7039 | -100.8667
5.7076 | -100.8402
5.7113 | -100.8175
5.7151 | -100.7954
5.7188 | -100.7878
5.7226 | -100.7942
5.7263 | -100.8146
5.7301 | -100.8421
5.7338 | -100.8789
5.7375 | -100.9213
5.7413 | -100.9615
5.745 | -100.9921
5.7488 | -101.0082
5.7525 | -101.0233
5.7563 | -101.0319
5.76 | -101.0307
5.7637 | -101.0257
5.7675 | -101.0236
5.7712 | -101.0236
5.775 | -101.0237
5.7787 | -101.0238
5.7825 | -101.023
5.7862 | -101.0277
5.7899 | -101.0303
5.7937 | -101.0436
5.7974 | -101.0574
5.8012 | -101.0789
5.8049 | -101.1013
5.8087 | -101.1281
5.8124 | -101.1585
5.8161 | -101.1968
5.8199 | -101.2433
5.8236 | -101.2929
5.8274 | -101.3554
5.8311 | -101.425
5.8349 | -101.5005
5.8386 | -101.582
5.8423 | -101.6671
5.8461 | -101.7634
5.8498 | -101.8438
5.8536 | -101.8955
5.8573 | -101.9329
5.8611 | -101.9514
5.8648 | -101.9407
5.8685 | -101.9141
5.8723 | -101.8691
5.876 | -101.8052
5.8798 | -101.725
5.8835 | -101.6466
5.8873 | -101.5682
5.891 | -101.4861
5.8947 | -101.4097
5.8985 | -101.3421
5.9022 | -101.2742
5.906 | -101.2086
5.9097 | -101.1569
5.9135 | -101.111
5.9172 | -101.0704
5.9209 | -101.0411
5.9247 | -101.0061
5.9284 | -100.9643
5.9322 | -100.9225
5.9359 | -100.8997
5.9396 | -100.8744
5.9434 | -100.8547
5.9471 | -100.8346
5.9509 | -100.812
5.9546 | -100.7967
5.9584 | -100.7824
5.9621 | -100.7757
5.9658 | -100.7765
5.9696 | -100.7758
5.9733 | -100.7713
5.9771 | -100.7671
5.9808 | -100.7531
5.9846 | -100.7272
5.9883 | -100.7135
5.992 | -100.7141
5.9958 | -100.7204
5.9995 | -100.7297
6.0033 | -100.7453
6.007 | -100.7759
6.0108 | -100.8126
6.0145 | -100.8607
6.0182 | -100.9181
6.022 | -100.9714
6.0257 | -101.0222
6.0295 | -101.0782
6.0332 | -101.1353
6.037 | -101.1995
6.0407 | -101.2617
6.0444 | -101.3175
6.0482 | -101.3574
6.0519 | -101.3881
6.0557 | -101.4018
6.0594 | -101.3983
6.0632 | -101.3809
6.0669 | -101.3553
6.0706 | -101.3174
6.0744 | -101.2781
6.0781 | -101.23
6.0819 | -101.167
6.0856 | -101.1052
6.0894 | -101.0539
6.0931 | -101.0176
6.0968 | -100.9996
6.1006 | -100.989
6.1043 | -100.9841
6.1081 | -100.9898
6.1118 | -101.0001
6.1156 | -101.0081
6.1193 | -101.0156
6.123 | -101.022
6.1268 | -101.0306
6.1305 | -101.0433
6.1343 | -101.0584
6.138 | -101.0798
6.1418 | -101.1083
6.1455 | -101.1334
6.1492 | -101.1685
6.153 | -101.2103
6.1567 | -101.2578
6.1605 | -101.3028
6.1642 | -101.3517
6.168 | -101.4022
6.1717 | -101.4521
6.1754 | -101.4891
6.1792 | -101.5186
6.1829 | -101.5387
6.1867 | -101.5375
6.1904 | -101.5188
6.1942 | -101.4917
6.1979 | -101.4625
6.2016 | -101.4197
6.2054 | -101.3761
6.2091 | -101.3397
6.2129 | -101.3049
6.2166 | -101.2832
6.2204 | -101.2794
6.2241 | -101.2896
6.2278 | -101.3203
6.2316 | -101.3607
6.2353 | -101.4033
6.2391 | -101.4523
6.2428 | -101.4988
6.2466 | -101.5428
6.2503 | -101.5702
6.254 | -101.5884
6.2578 | -101.5986
6.2615 | -101.5928
6.2653 | -101.5739
6.269 | -101.5528
6.2727 | -101.528
6.2765 | -101.5001
6.2802 | -101.4814
6.284 | -101.4702
6.2877 | -101.4631
6.2915 | -101.4624
6.2952 | -101.477
6.2989 | -101.499
6.3027 | -101.5298
6.3064 | -101.5651
6.3102 | -101.5982
6.3139 | -101.6276
6.3177 | -101.6502
6.3214 | -101.6677
6.3251 | -101.6854
6.3289 | -101.6993
6.3326 | -101.7039
6.3364 | -101.6988
6.3401 | -101.6864
6.3439 | -101.6789
6.3476 | -101.6716
6.3513 | -101.6584
6.3551 | -101.6521
6.3588 | -101.6536
6.3626 | -101.6535
6.3663 | -101.647
6.3701 | -101.6413
6.3738 | -101.6369
6.3775 | -101.6271
6.3813 | -101.6074
6.385 | -101.5828
6.3888 | -101.5464
6.3925 | -101.5046
6.3963 | -101.4619
6.4 | -101.4228
6.4037 | -101.3879
6.4075 | -101.3576
6.4112 | -101.3413
6.415 | -101.3388
6.4187 | -101.3484
6.4225 | -101.3645
6.4262 | -101.3854
6.4299 | -101.4159
6.4337 | -101.4562
6.4374 | -101.4977
6.4412 | -101.5408
6.4449 | -101.5898
6.4487 | -101.6395
6.4524 | -101.6829
6.4561 | -101.7211
6.4599 | -101.7582
6.4636 | -101.7844
6.4674 | -101.7996
6.4711 | -101.8124
6.4749 | -101.8239
6.4786 | -101.8262
6.4823 | -101.8341
6.4861 | -101.8424
6.4898 | -101.8472
6.4936 | -101.8392
6.4973 | -101.8285
6.5011 | -101.8143
6.5048 | -101.7958
6.5085 | -101.7674
6.5123 | -101.7324
6.516 | -101.6985
6.5198 | -101.6725
6.5235 | -101.6523
6.5273 | -101.6357
6.531 | -101.6182
6.5347 | -101.6044
6.5385 | -101.5968
6.5422 | -101.5887
6.546 | -101.5802
6.5497 | -101.5711
6.5534 | -101.5657
6.5572 | -101.5629
6.5609 | -101.5634
6.5647 | -101.5601
6.5684 | -101.5545
6.5722 | -101.5554
6.5759 | -101.5621
6.5796 | -101.5611
6.5834 | -101.5621
6.5871 | -101.577
6.5909 | -101.5984
6.5946 | -101.6328
6.5984 | -101.6678
6.6021 | -101.6986
6.6058 | -101.7313
6.6096 | -101.7617
6.6133 | -101.7857
6.6171 | -101.7966
6.6208 | -101.8005
6.6246 | -101.8101
6.6283 | -101.8215
6.632 | -101.8371
6.6358 | -101.8573
6.6395 | -101.8732
6.6433 | -101.8862
6.647 | -101.9066
6.6508 | -101.9335
6.6545 | -101.9558
6.6582 | -101.9751
6.662 | -101.9933
6.6657 | -102.0065
6.6695 | -102.0073
6.6732 | -101.9958
6.677 | -101.9806
6.6807 | -101.9692
6.6844 | -101.9477
6.6882 | -101.9086
6.6919 | -101.8656
6.6957 | -101.8189
6.6994 | -101.7768
6.7032 | -101.7402
6.7069 | -101.7045
6.7106 | -101.6832
6.7144 | -101.6791
6.7181 | -101.6923
6.7219 | -101.7137
6.7256 | -101.7422
6.7294 | -101.7736
6.7331 | -101.8197
6.7368 | -101.8699
6.7406 | -101.9245
6.7443 | -101.976
6.7481 | -102.0106
6.7518 | -102.0341
6.7556 | -102.0539
6.7593 | -102.0713
6.763 | -102.0826
6.7668 | -102.088
6.7705 | -102.0916
6.7743 | -102.0897
6.778 | -102.0818
6.7818 | -102.0661
6.7855 | -102.0524
6.7892 | -102.0429
6.793 | -102.028
6.7967 | -102.0095
6.8005 | -101.9898
6.8042 | -101.9754
6.808 | -101.9606
6.8117 | -101.9439
6.8154 | -101.9293
6.8192 | -101.9184
6.8229 | -101.921
6.8267 | -101.9274
6.8304 | -101.9427
6.8342 | -101.9584
6.8379 | -101.9806
6.8416 | -102.0121
6.8454 | -102.0536
6.8491 | -102.0951
6.8529 | -102.1356
6.8566 | -102.1732
6.8604 | -102.2109
6.8641 | -102.2519
6.8678 | -102.296
6.8716 | -102.3296
6.8753 | -102.3549
6.8791 | -102.368
6.8828 | -102.3741
6.8865 | -102.3679
6.8903 | -102.3545
6.894 | -102.34
6.8978 | -102.3094
6.9015 | -102.2716
6.9053 | -102.2318
6.909 | -102.1909
6.9127 | -102.1466
6.9165 | -102.1034
6.9202 | -102.0665
6.924 | -102.035
6.9277 | -102.0044
6.9315 | -101.9824
6.9352 | -101.9712
6.9389 | -101.9694
6.9427 | -101.9747
6.9464 | -101.9879
6.9502 | -102.0125
6.9539 | -102.0436
6.9577 | -102.087
6.9614 | -102.1338
6.9651 | -102.1837
6.9689 | -102.241
6.9726 | -102.3034
6.9764 | -102.3623
6.9801 | -102.424
6.9839 | -102.4823
6.9876 | -102.5254
6.9913 | -102.5547
6.9951 | -102.5715
6.9988 | -102.5858
7.0026 | -102.575
7.0063 | -102.5443
7.0101 | -102.5117
7.0138 | -102.4709
7.0175 | -102.412
7.0213 | -102.3633
7.025 | -102.3156
7.0288 | -102.2652
7.0325 | -102.2216
7.0363 | -102.1846
7.04 | -102.1518
7.0437 | -102.1279
7.0475 | -102.109
7.0512 | -102.0992
7.055 | -102.095
7.0587 | -102.0991
7.0625 | -102.0994
7.0662 | -102.105
7.0699 | -102.1105
7.0737 | -102.1105
7.0774 | -102.1156
7.0812 | -102.1208
7.0849 | -102.12
7.0887 | -102.1256
7.0924 | -102.1436
7.0961 | -102.1648
7.0999 | -102.1902
7.1036 | -102.2273
7.1074 | -102.2745
7.1111 | -102.3259
7.1149 | -102.3775
7.1186 | -102.4291
7.1223 | -102.4774
7.1261 | -102.5194
7.1298 | -102.5589
7.1336 | -102.5858
7.1373 | -102.6022
7.1411 | -102.6082
7.1448 | -102.5978
7.1485 | -102.5726
7.1523 | -102.5384
7.156 | -102.5044
7.1598 | -102.4659
7.1635 | -102.4263
7.1673 | -102.3903
7.171 | -102.3705
7.1747 | -102.3623
7.1785 | -102.3675
7.1822 | -102.3859
7.186 | -102.4121
7.1897 | -102.4474
7.1935 | -102.4909
7.1972 | -102.5382
7.2009 | -102.5821
7.2047 | -102.6271
7.2084 | -102.6771
7.2122 | -102.7215
7.2159 | -102.7614
7.2196 | -102.7966
7.2234 | -102.8273
7.2271 | -102.8545
7.2309 | -102.8747
7.2346 | -102.8867
7.2384 | -102.8814
7.2421 | -102.8628
7.2458 | -102.8364
7.2496 | -102.7987
7.2533 | -102.748
7.2571 | -102.6936
7.2608 | -102.6372
7.2646 | -102.5764
7.2683 | -102.5081
7.272 | -102.4414
7.2758 | -102.3731
7.2795 | -102.3065
7.2833 | -102.2455
7.287 | -102.1941
7.2908 | -102.1518
7.2945 | -102.1194
7.2982 | -102.0967
7.302 | -102.0855
7.3057 | -102.09
7.3095 | -102.1108
7.3132 | -102.1402
7.317 | -102.177
7.3207 | -102.2201
7.3244 | -102.2712
7.3282 | -102.3193
7.3319 | -102.3639
7.3357 | -102.4031
7.3394 | -102.4386
7.3432 | -102.4695
7.3469 | -102.4981
7.3506 | -102.5188
7.3544 | -102.5361
7.3581 | -102.5537
7.3619 | -102.5749
7.3656 | -102.594
7.3694 | -102.613
7.3731 | -102.6302
7.3768 | -102.6581
7.3806 | -102.6847
7.3843 | -102.7131
7.3881 | -102.7418
7.3918 | -102.7683
7.3956 | -102.7926
7.3993 | -102.813
7.403 | -102.8332
7.4068 | -102.8534
7.4105 | -102.8678
7.4143 | -102.8703
7.418 | -102.8724
7.4218 | -102.874
7.4255 | -102.8754
7.4292 | -102.8706
7.433 | -102.8629
7.4367 | -102.8535
7.4405 | -102.8389
7.4442 | -102.8277
7.448 | -102.8129
7.4517 | -102.7983
7.4554 | -102.7824
7.4592 | -102.7718
7.4629 | -102.7614
7.4667 | -102.756
7.4704 | -102.7483
7.4742 | -102.7408
7.4779 | -102.7421
7.4816 | -102.7566
7.4854 | -102.7737
7.4891 | -102.7896
7.4929 | -102.8061
7.4966 | -102.8214
7.5004 | -102.8293
7.5041 | -102.8267
7.5078 | -102.8191
7.5116 | -102.7966
7.5153 | -102.7646
7.5191 | -102.7257
7.5228 | -102.6777
7.5266 | -102.6223
7.5303 | -102.5644
7.534 | -102.5021
7.5378 | -102.4419
7.5415 | -102.3887
7.5453 | -102.3399
7.549 | -102.2969
7.5527 | -102.2624
7.5565 | -102.2785
7.5602 | -102.273
7.564 | -102.2783
7.5677 | -102.2985
7.5715 | -102.3277
7.5752 | -102.3666
7.5789 | -102.4142
7.5827 | -102.4733
7.5864 | -102.539
7.5902 | -102.5998
7.5939 | -102.6603
7.5977 | -102.7192
7.6014 | -102.7786
7.6051 | -102.8304
7.6089 | -102.8741
7.6126 | -102.9089
7.6164 | -102.9387
7.6201 | -102.9484
7.6239 | -102.9448
7.6276 | -102.9299
7.6313 | -102.9043
7.6351 | -102.8669
7.6388 | -102.8247
7.6426 | -102.7795
7.6463 | -102.7302
7.6501 | -102.6846
7.6538 | -102.6395
7.6575 | -102.6026
7.6613 | -102.5742
7.665 | -102.5528
7.6688 | -102.5393
7.6725 | -102.5374
7.6763 | -102.5471
7.68 | -102.5708
7.6837 | -102.5991
7.6875 | -102.6319
7.6912 | -102.6723
7.695 | -102.7163
7.6987 | -102.7646
7.7025 | -102.8128
7.7062 | -102.8539
7.7099 | -102.8922
7.7137 | -102.9282
7.7174 | -102.9509
7.7212 | -102.9644
7.7249 | -102.9754
7.7287 | -102.9807
7.7324 | -102.969
7.7361 | -102.9503
7.7399 | -102.9296
7.7436 | -102.9058
7.7474 | -102.8811
7.7511 | -102.8553
7.7549 | -102.8211
7.7586 | -102.7855
7.7623 | -102.7505
7.7661 | -102.7077
7.7698 | -102.666
7.7736 | -102.6257
7.7773 | -102.5807
7.7811 | -102.5445
7.7848 | -102.5127
7.7885 | -102.4842
7.7923 | -102.4631
7.796 | -102.4498
7.7998 | -102.4482
7.8035 | -102.4604
7.8073 | -102.4785
7.811 | -102.506
7.8147 | -102.5436
7.8185 | -102.5906
7.8222 | -102.6514
7.826 | -102.7114
7.8297 | -102.7734
7.8335 | -102.8393
7.8372 | -102.8989
7.8409 | -102.9494
7.8447 | -102.9965
7.8484 | -103.0408
7.8522 | -103.0721
7.8559 | -103.0894
7.8596 | -103.0941
7.8634 | -103.0899
7.8671 | -103.0831
7.8709 | -103.0748
7.8746 | -103.0558
7.8784 | -103.0324
7.8821 | -103.0116
7.8858 | -102.9815
7.8896 | -102.9425
7.8933 | -102.9114
7.8971 | -102.8747
7.9008 | -102.8391
7.9046 | -102.8079
7.9083 | -102.7746
7.912 | -102.7412
7.9158 | -102.7156
7.9195 | -102.6927
7.9233 | -102.6749
7.927 | -102.6649
7.9308 | -102.6601
7.9345 | -102.6659
7.9382 | -102.6763
7.942 | -102.6998
7.9457 | -102.7325
7.9495 | -102.7722
7.9532 | -102.8129
7.957 | -102.8543
7.9607 | -102.8947
7.9644 | -102.9309
7.9682 | -102.9704
7.9719 | -103.0083
7.9757 | -103.0403
7.9794 | -103.059
7.9832 | -103.0724
7.9869 | -103.0719
7.9906 | -103.0683
7.9944 | -103.0601
7.9981 | -103.0404
8.0019 | -103.0207
8.0056 | -103.0006
8.0094 | -102.9747
8.0131 | -102.9492
8.0168 | -102.9282
8.0206 | -102.911
8.0243 | -102.8998
8.0281 | -102.895
8.0318 | -102.8897
8.0356 | -102.8832
8.0393 | -102.8827
8.043 | -102.8841
8.0468 | -102.8851
8.0505 | -102.8857
8.0543 | -102.8913
8.058 | -102.8977
8.0618 | -102.9002
8.0655 | -102.9033
8.0692 | -102.9117
8.073 | -102.9201
8.0767 | -102.9319
8.0805 | -102.9533
8.0842 | -102.9741
8.088 | -102.9911
8.0917 | -103.014
8.0954 | -103.0439
8.0992 | -103.0666
8.1029 | -103.0888
8.1067 | -103.1085
8.1104 | -103.1237
8.1142 | -103.1369
8.1179 | -103.1502
8.1216 | -103.1596
8.1254 | -103.1706
8.1291 | -103.1838
8.1329 | -103.1931
8.1366 | -103.1973
8.1404 | -103.1985
8.1441 | -103.1968
8.1478 | -103.1917
8.1516 | -103.1918
8.1553 | -103.1925
8.1591 | -103.1781
8.1628 | -103.1626
8.1666 | -103.1463
8.1703 | -103.133
8.174 | -103.1161
8.1778 | -103.0977
8.1815 | -103.0874
8.1853 | -103.0808
8.189 | -103.0715
8.1927 | -103.062
8.1965 | -103.05
8.2002 | -103.037
8.204 | -103.0262
8.2077 | -103.0152
8.2115 | -103.0004
8.2152 | -102.9911
8.2189 | -102.9877
8.2227 | -102.9914
8.2264 | -102.9938
8.2302 | -102.9992
8.2339 | -103.0037
8.2377 | -103.0067
8.2414 | -103.0134
8.2451 | -103.0223
8.2489 | -103.025
8.2526 | -103.024
8.2564 | -103.0282
8.2601 | -103.0359
8.2639 | -103.0433
8.2676 | -103.0493
8.2713 | -103.0492
8.2751 | -103.0418
8.2788 | -103.0357
8.2826 | -103.0376
8.2863 | -103.035
8.2901 | -103.0341
8.2938 | -103.0334
8.2975 | -103.0363
8.3013 | -103.0439
8.305 | -103.0526
8.3088 | -103.0635
8.3125 | -103.0856
8.3163 | -103.1149
8.32 | -103.1499
8.3237 | -103.1889
8.3275 | -103.2309
8.3312 | -103.2779
8.335 | -103.3278
8.3387 | -103.3764
8.3425 | -103.4219
8.3462 | -103.4634
8.3499 | -103.5016
8.3537 | -103.5349
8.3574 | -103.5584
8.3612 | -103.5754
8.3649 | -103.5854
8.3687 | -103.5873
8.3724 | -103.5824
8.3761 | -103.5716
8.3799 | -103.5528
8.3836 | -103.5276
8.3874 | -103.4991
8.3911 | -103.463
8.3949 | -103.4282
8.3986 | -103.3965
8.4023 | -103.3638
8.4061 | -103.3284
8.4098 | -103.2913
8.4136 | -103.2548
8.4173 | -103.2194
8.4211 | -103.1894
8.4248 | -103.1594
8.4285 | -103.1318
8.4323 | -103.1102
8.436 | -103.0921
8.4398 | -103.0713
8.4435 | -103.0626
8.4473 | -103.06
8.451 | -103.0626
8.4547 | -103.0703
8.4585 | -103.0753
8.4622 | -103.0815
8.466 | -103.0974
8.4697 | -103.1166
8.4734 | -103.1345
8.4772 | -103.152
8.4809 | -103.1663
8.4847 | -103.1825
8.4884 | -103.1972
8.4922 | -103.209
8.4959 | -103.2211
8.4996 | -103.2311
8.5034 | -103.2411
8.5071 | -103.248
8.5109 | -103.249
8.5146 | -103.2518
8.5184 | -103.2527
8.5221 | -103.2545
8.5258 | -103.2571
8.5296 | -103.26
8.5333 | -103.2654
8.5371 | -103.2702
8.5408 | -103.2728
8.5446 | -103.2766
8.5483 | -103.282
8.552 | -103.2845
8.5558 | -103.2881
8.5595 | -103.2911
8.5633 | -103.3032
8.567 | -103.3067
8.5708 | -103.3109
8.5745 | -103.3246
8.5782 | -103.3386
8.582 | -103.3565
8.5857 | -103.3745
8.5895 | -103.3866
8.5932 | -103.3972
8.597 | -103.4098
8.6007 | -103.4184
8.6044 | -103.4285
8.6082 | -103.4254
8.6119 | -103.4216
8.6157 | -103.4165
8.6194 | -103.4119
8.6232 | -103.4029
8.6269 | -103.3905
8.6306 | -103.3785
8.6344 | -103.3704
8.6381 | -103.3601
8.6419 | -103.3534
8.6456 | -103.3478
8.6494 | -103.3375
8.6531 | -103.3318
8.6568 | -103.3331
8.6606 | -103.3395
8.6643 | -103.3461
8.6681 | -103.3512
8.6718 | -103.3596
8.6756 | -103.3719
8.6793 | -103.3861
8.683 | -103.3951
8.6868 | -103.4024
8.6905 | -103.4072
8.6943 | -103.4164
8.698 | -103.4203
8.7018 | -103.4179
8.7055 | -103.4142
8.7092 | -103.4053
8.713 | -103.3901
8.7167 | -103.3763
8.7205 | -103.3599
8.7242 | -103.3411
8.728 | -103.3227
8.7317 | -103.3028
8.7354 | -103.2863
8.7392 | -103.2716
8.7429 | -103.2633
8.7467 | -103.2537
8.7504 | -103.2484
8.7542 | -103.2555
8.7579 | -103.2663
8.7616 | -103.2778
8.7654 | -103.2943
8.7691 | -103.3189
8.7729 | -103.3481
8.7766 | -103.3766
8.7804 | -103.4092
8.7841 | -103.4359
8.7878 | -103.4603
8.7916 | -103.4845
8.7953 | -103.5107
8.7991 | -103.5398
8.8028 | -103.5607
8.8065 | -103.5776
8.8103 | -103.5946
8.814 | -103.6135
8.8178 | -103.6316
8.8215 | -103.6437
8.8253 | -103.6486
8.829 | -103.6532
8.8327 | -103.6537
8.8365 | -103.65
8.8402 | -103.6458
8.844 | -103.6378
8.8477 | -103.6248
8.8515 | -103.6121
8.8552 | -103.5958
8.8589 | -103.5755
8.8627 | -103.5556
8.8664 | -103.5369
8.8702 | -103.5182
8.8739 | -103.4998
8.8777 | -103.4882
8.8814 | -103.478
8.8851 | -103.4647
8.8889 | -103.4569
8.8926 | -103.4516
8.8964 | -103.4482
8.9001 | -103.443
8.9039 | -103.4437
8.9076 | -103.4489
8.9113 | -103.4574
8.9151 | -103.4732
8.9188 | -103.4925
8.9226 | -103.5133
8.9263 | -103.5369
8.9301 | -103.5605
8.9338 | -103.588
8.9375 | -103.6126
8.9413 | -103.6348
8.945 | -103.6551
8.9488 | -103.6681
8.9525 | -103.6844
8.9563 | -103.7021
8.96 | -103.7131
8.9637 | -103.7199
8.9675 | -103.7251
8.9712 | -103.7279
8.975 | -103.731
8.9787 | -103.732
8.9825 | -103.7264
8.9862 | -103.7216
8.9899 | -103.7185
8.9937 | -103.7063
8.9974 | -103.6878
9.0012 | -103.6714
9.0049 | -103.6535
9.0087 | -103.6357
9.0124 | -103.6188
9.0161 | -103.5988
9.0199 | -103.5807
9.0236 | -103.5651
9.0274 | -103.5492
9.0311 | -103.5402
9.0349 | -103.5285
9.0386 | -103.5208
9.0423 | -103.518
9.0461 | -103.5207
9.0498 | -103.5258
9.0536 | -103.5377
9.0573 | -103.5544
9.0611 | -103.5756
9.0648 | -103.603
9.0685 | -103.6328
9.0723 | -103.6664
9.076 | -103.6996
9.0798 | -103.7382
9.0835 | -103.7746
9.0873 | -103.8063
9.091 | -103.8411
9.0947 | -103.8701
9.0985 | -103.894
9.1022 | -103.9084
9.106 | -103.9108
9.1097 | -103.9106
9.1135 | -103.9087
9.1172 | -103.8985
9.1209 | -103.8806
9.1247 | -103.8655
9.1284 | -103.8467
9.1322 | -103.8166
9.1359 | -103.7862
9.1396 | -103.7482
9.1434 | -103.7086
9.1471 | -103.6751
9.1509 | -103.6375
9.1546 | -103.6004
9.1584 | -103.5645
9.1621 | -103.5261
9.1658 | -103.4953
9.1696 | -103.4699
9.1733 | -103.4486
9.1771 | -103.4495
9.1808 | -103.4406
9.1846 | -103.4338
9.1883 | -103.435
9.192 | -103.4458
9.1958 | -103.4651
9.1995 | -103.4895
9.2033 | -103.5254
9.207 | -103.5654
9.2108 | -103.6118
9.2145 | -103.6617
9.2182 | -103.7131
9.222 | -103.7685
9.2257 | -103.82
9.2295 | -103.8725
9.2332 | -103.9243
9.237 | -103.9762
9.2407 | -104.0273
9.2444 | -104.0749
9.2482 | -104.1127
9.2519 | -104.1375
9.2557 | -104.1532
9.2594 | -104.1552
9.2632 | -104.1484
9.2669 | -104.1336
9.2706 | -104.1131
9.2744 | -104.0871
9.2781 | -104.0536
9.2819 | -104.0124
9.2856 | -103.9753
9.2894 | -103.9404
9.2931 | -103.9076
9.2968 | -103.8764
9.3006 | -103.8522
9.3043 | -103.8332
9.3081 | -103.8216
9.3118 | -103.8177
9.3156 | -103.8204
9.3193 | -103.8282
9.323 | -103.8473
9.3268 | -103.8727
9.3305 | -103.8986
9.3343 | -103.9279
9.338 | -103.9617
9.3418 | -103.9956
9.3455 | -104.028
9.3492 | -104.0619
9.353 | -104.0948
9.3567 | -104.1229
9.3605 | -104.1454
9.3642 | -104.1697
9.368 | -104.1891
9.3717 | -104.1992
9.3754 | -104.2016
9.3792 | -104.196
9.3829 | -104.1848
9.3867 | -104.1716
9.3904 | -104.1583
9.3942 | -104.1466
9.3979 | -104.1352
9.4016 | -104.121
9.4054 | -104.1062
9.4091 | -104.0904
9.4129 | -104.0768
9.4166 | -104.0648
9.4204 | -104.0556
9.4241 | -104.0439
9.4278 | -104.0331
9.4316 | -104.0261
9.4353 | -104.0251
9.4391 | -104.0245
9.4428 | -104.024
9.4466 | -104.0252
9.4503 | -104.0265
9.454 | -104.035
9.4578 | -104.0436
9.4615 | -104.0544
9.4653 | -104.0641
9.469 | -104.0801
9.4727 | -104.0957
9.4765 | -104.1106
9.4802 | -104.1308
9.484 | -104.1536
9.4877 | -104.1725
9.4915 | -104.1927
9.4952 | -104.216
9.4989 | -104.2359
9.5027 | -104.2485
9.5064 | -104.2615
9.5102 | -104.2761
9.5139 | -104.2914
9.5177 | -104.3027
9.5214 | -104.3041
9.5251 | -104.3068
9.5289 | -104.3028
9.5326 | -104.2862
9.5364 | -104.2721
9.5401 | -104.2534
9.5439 | -104.2275
9.5476 | -104.206
9.5513 | -104.1837
9.5551 | -104.1574
9.5588 | -104.1331
9.5626 | -104.1077
9.5663 | -104.0836
9.5701 | -104.058
9.5738 | -104.0317
9.5775 | -104.0022
9.5813 | -103.9737
9.585 | -103.9496
9.5888 | -103.9294
9.5925 | -103.9091
9.5963 | -103.8924
9.6 | -103.8747
9.6037 | -103.8678
9.6075 | -103.8669
9.6112 | -103.8699
9.615 | -103.8834
9.6187 | -103.9022
9.6225 | -103.9272
9.6262 | -103.9592
9.6299 | -103.9944
9.6337 | -104.0335
9.6374 | -104.0762
9.6412 | -104.121
9.6449 | -104.1616
9.6487 | -104.1945
9.6524 | -104.2269
9.6561 | -104.2543
9.6599 | -104.2727
9.6636 | -104.287
9.6674 | -104.2976
9.6711 | -104.3026
9.6749 | -104.2917
9.6786 | -104.2812
9.6823 | -104.2659
9.6861 | -104.2362
9.6898 | -104.2036
9.6936 | -104.1718
9.6973 | -104.1343
9.7011 | -104.0978
9.7048 | -104.0565
9.7085 | -104.0166
9.7123 | -103.9788
9.716 | -103.9574
9.7198 | -103.9329
9.7235 | -103.9109
9.7273 | -103.8889
9.731 | -103.8748
9.7347 | -103.8646
9.7385 | -103.8571
9.7422 | -103.8543
9.746 | -103.8544
9.7497 | -103.8594
9.7535 | -103.8718
9.7572 | -103.8802
9.7609 | -103.8848
9.7647 | -103.893
9.7684 | -103.9061
9.7722 | -103.9174
9.7759 | -103.926
9.7796 | -103.9322
9.7834 | -103.9367
9.7871 | -103.9463
9.7909 | -103.9566
9.7946 | -103.9629
9.7984 | -103.9738
9.8021 | -103.9883
9.8058 | -103.9949
9.8096 | -103.9996
9.8133 | -104.005
9.8171 | -104.0112
9.8208 | -104.018
9.8246 | -104.0272
9.8283 | -104.0342
9.832 | -104.0465
9.8358 | -104.0574
9.8395 | -104.0673
9.8433 | -104.0784
9.847 | -104.0919
9.8508 | -104.1074
9.8545 | -104.127
9.8582 | -104.1438
9.862 | -104.1579
9.8657 | -104.1694
9.8695 | -104.1769
9.8732 | -104.1795
9.877 | -104.1769
9.8807 | -104.1754
9.8844 | -104.1745
9.8882 | -104.1725
9.8919 | -104.162
9.8957 | -104.1427
9.8994 | -104.1216
9.9032 | -104.104
9.9069 | -104.0901
9.9106 | -104.0736
9.9144 | -104.0528
9.9181 | -104.0328
9.9219 | -104.014
9.9256 | -103.9991
9.9294 | -103.986
9.9331 | -103.9703
9.9368 | -103.9593
9.9406 | -103.948
9.9443 | -103.9409
9.9481 | -103.9361
9.9518 | -103.938
9.9556 | -103.9455
9.9593 | -103.9554
9.963 | -103.9698
9.9668 | -103.9819
9.9705 | -103.9956
9.9743 | -104.0173
9.978 | -104.0364
9.9818 | -104.0544
9.9855 | -104.0762
9.9892 | -104.0982
9.993 | -104.12
9.9967 | -104.1383
10.0005 | -104.1518
10.0042 | -104.1666
10.008 | -104.1804
10.0117 | -104.1952
10.0154 | -104.2104
10.0192 | -104.2238
10.0229 | -104.23
10.0267 | -104.2285
10.0304 | -104.2285
10.0342 | -104.2293
10.0379 | -104.2281
10.0416 | -104.2231
10.0454 | -104.216
10.0491 | -104.2066
10.0529 | -104.1936
10.0566 | -104.1823
10.0604 | -104.1719
10.0641 | -104.1611
10.0678 | -104.1523
10.0716 | -104.1385
10.0753 | -104.1213
10.0791 | -104.1017
10.0828 | -104.0794
10.0866 | -104.0591
10.0903 | -104.0381
10.094 | -104.0176
10.0978 | -103.995
10.1015 | -103.9707
10.1053 | -103.9491
10.109 | -103.9318
10.1127 | -103.9227
10.1165 | -103.9127
10.1202 | -103.9029
10.124 | -103.8977
10.1277 | -103.8987
10.1315 | -103.8972
10.1352 | -103.9024
10.1389 | -103.9092
10.1427 | -103.9182
10.1464 | -103.9284
10.1502 | -103.945
10.1539 | -103.9627
10.1577 | -103.9783
10.1614 | -103.9981
10.1651 | -104.0172
10.1689 | -104.0384
10.1726 | -104.0595
10.1764 | -104.0725
10.1801 | -104.0838
10.1839 | -104.0964
10.1876 | -104.1133
10.1913 | -104.1201
10.1951 | -104.1213
10.1988 | -104.1183
10.2026 | -104.1092
10.2063 | -104.0991
10.2101 | -104.0867
10.2138 | -104.0696
10.2175 | -104.0568
10.2213 | -104.0432
10.225 | -104.0309
10.2288 | -104.02
10.2325 | -104.0122
10.2363 | -104.0066
10.24 | -104.0056
10.2437 | -104.0086
10.2475 | -104.0112
10.2512 | -104.0167
10.255 | -104.0257
10.2587 | -104.0443
10.2625 | -104.0541
10.2662 | -104.0679
10.2699 | -104.0802
10.2737 | -104.0931
10.2774 | -104.103
10.2812 | -104.1129
10.2849 | -104.1181
10.2887 | -104.1221
10.2924 | -104.1251
10.2961 | -104.126
10.2999 | -104.1187
10.3036 | -104.1126
10.3074 | -104.1098
10.3111 | -104.1043
10.3149 | -104.0998
10.3186 | -104.0909
10.3223 | -104.0857
10.3261 | -104.0854
10.3298 | -104.0859
10.3336 | -104.089
10.3373 | -104.0917
10.3411 | -104.1013
10.3448 | -104.1123
10.3485 | -104.1272
10.3523 | -104.1483
10.356 | -104.1719
10.3598 | -104.1982
10.3635 | -104.2267
10.3673 | -104.2558
10.371 | -104.29
10.3747 | -104.322
10.3785 | -104.3504
10.3822 | -104.3767
10.386 | -104.4021
10.3897 | -104.4232
10.3934 | -104.4381
10.3972 | -104.4531
10.4009 | -104.4632
10.4047 | -104.4676
10.4084 | -104.4685
10.4122 | -104.4651
10.4159 | -104.4565
10.4196 | -104.4473
10.4234 | -104.4402
10.4271 | -104.4304
10.4309 | -104.4227
10.4346 | -104.4139
10.4384 | -104.4056
10.4421 | -104.4004
10.4458 | -104.3922
10.4496 | -104.3856
10.4533 | -104.3848
10.4571 | -104.3819
10.4608 | -104.3754
10.4646 | -104.3688
10.4683 | -104.3609
10.472 | -104.3561
10.4758 | -104.3473
10.4795 | -104.3352
10.4833 | -104.3198
10.487 | -104.3022
10.4908 | -104.2841
10.4945 | -104.2642
10.4982 | -104.242
10.502 | -104.2189
10.5057 | -104.1945
10.5095 | -104.1665
10.5132 | -104.1412
10.517 | -104.1215
10.5207 | -104.1043
10.5244 | -104.0892
10.5282 | -104.0742
10.5319 | -104.0657
10.5357 | -104.0594
10.5394 | -104.0553
10.5432 | -104.0568
10.5469 | -104.0563
10.5506 | -104.0542
10.5544 | -104.053
10.5581 | -104.0514
10.5619 | -104.0471
10.5656 | -104.0409
10.5694 | -104.0383
10.5731 | -104.0345
10.5768 | -104.0269
10.5806 | -104.0202
10.5843 | -104.0079
10.5881 | -103.9942
10.5918 | -103.9782
10.5956 | -103.9621
10.5993 | -103.9488
10.603 | -103.9357
10.6068 | -103.9197
10.6105 | -103.9038
10.6143 | -103.8877
10.618 | -103.8765
10.6218 | -103.8693
10.6255 | -103.8604
10.6292 | -103.8549
10.633 | -103.8502
10.6367 | -103.8464
10.6405 | -103.848
10.6442 | -103.8535
10.648 | -103.8578
10.6517 | -103.8663
10.6554 | -103.8776
10.6592 | -103.8882
10.6629 | -103.9003
10.6667 | -103.9113
10.6704 | -103.9218
10.6742 | -103.9325
10.6779 | -103.9385
10.6816 | -103.9491
10.6854 | -103.9582
10.6891 | -103.9631
10.6929 | -103.9623
10.6966 | -103.9573
10.7004 | -103.9546
10.7041 | -103.9535
10.7078 | -103.9495
10.7116 | -103.9458
10.7153 | -103.9431
10.7191 | -103.9377
10.7228 | -103.9263
10.7265 | -103.9119
10.7303 | -103.8981
10.734 | -103.884
10.7378 | -103.8716
10.7415 | -103.8593
10.7453 | -103.8466
10.749 | -103.8257
10.7527 | -103.8088
10.7565 | -103.7933
10.7602 | -103.7789
10.764 | -103.7673
10.7677 | -103.759
10.7715 | -103.7482
10.7752 | -103.7351
10.7789 | -103.7261
10.7827 | -103.7153
10.7864 | -103.7059
10.7902 | -103.6966
10.7939 | -103.6899
10.7977 | -103.6898
10.8014 | -103.6854
10.8051 | -103.685
10.8089 | -103.6868
10.8126 | -103.6923
10.8164 | -103.699
10.8201 | -103.7133
10.8239 | -103.7297
10.8276 | -103.745
10.8313 | -103.766
10.8351 | -103.7968
10.8388 | -103.8278
10.8426 | -103.8584
10.8463 | -103.8913
10.8501 | -103.9272
10.8538 | -103.9654
10.8575 | -104
10.8613 | -104.0279
10.865 | -104.0514
10.8688 | -104.0706
10.8725 | -104.0871
10.8763 | -104.1005
10.88 | -104.1122
10.8837 | -104.1204
10.8875 | -104.1239
10.8912 | -104.123
10.895 | -104.1241
10.8987 | -104.1224
10.9025 | -104.1139
10.9062 | -104.1061
10.9099 | -104.0972
10.9137 | -104.0892
10.9174 | -104.0814
10.9212 | -104.0755
10.9249 | -104.068
10.9287 | -104.0641
10.9324 | -104.0597
10.9361 | -104.0522
10.9399 | -104.0491
10.9436 | -104.049
10.9474 | -104.0515
10.9511 | -104.05
10.9549 | -104.0532
10.9586 | -104.056
10.9623 | -104.0646
10.9661 | -104.0718
10.9698 | -104.0768
10.9736 | -104.0861
10.9773 | -104.1006
10.9811 | -104.1111
10.9848 | -104.1274
10.9885 | -104.1462
10.9923 | -104.1614
10.996 | -104.1737
10.9998 | -104.1886
11.0035 | -104.2086
11.0073 | -104.2286
11.011 | -104.248
11.0147 | -104.2654
11.0185 | -104.2824
11.0222 | -104.2954
11.026 | -104.305
11.0297 | -104.3141
11.0335 | -104.3231
11.0372 | -104.3303
11.0409 | -104.3334
11.0447 | -104.3308
11.0484 | -104.325
11.0522 | -104.3164
11.0559 | -104.306
11.0596 | -104.2977
11.0634 | -104.2926
11.0671 | -104.2895
11.0709 | -104.2772
11.0746 | -104.2568
11.0784 | -104.2423
11.0821 | -104.2336
11.0858 | -104.2294
11.0896 | -104.2254
11.0933 | -104.2253
11.0971 | -104.2294
11.1008 | -104.2313
11.1046 | -104.2348
11.1083 | -104.2418
11.112 | -104.2507
11.1158 | -104.2592
11.1195 | -104.2698
11.1233 | -104.2805
11.127 | -104.2944
11.1308 | -104.3083
11.1345 | -104.3151
11.1382 | -104.3243
11.142 | -104.3322
11.1457 | -104.3323
11.1495 | -104.3314
11.1532 | -104.3315
11.157 | -104.3282
11.1607 | -104.3235
11.1644 | -104.3094
11.1682 | -104.2946
11.1719 | -104.2748
11.1757 | -104.2556
11.1794 | -104.2364
11.1832 | -104.2175
11.1869 | -104.1981
11.1906 | -104.171
11.1944 | -104.1415
11.1981 | -104.115
11.2019 | -104.0935
11.2056 | -104.0773
11.2094 | -104.0632
11.2131 | -104.0532
11.2168 | -104.0434
11.2206 | -104.0354
11.2243 | -104.0319
11.2281 | -104.0333
11.2318 | -104.0402
11.2356 | -104.0557
11.2393 | -104.0707
11.243 | -104.0827
11.2468 | -104.0961
11.2505 | -104.1127
11.2543 | -104.1286
11.258 | -104.1417
11.2618 | -104.1514
11.2655 | -104.1599
11.2692 | -104.167
11.273 | -104.1723
11.2767 | -104.1758
11.2805 | -104.1728
11.2842 | -104.1672
11.288 | -104.1578
11.2917 | -104.1483
11.2954 | -104.1393
11.2992 | -104.126
11.3029 | -104.1106
11.3067 | -104.0987
11.3104 | -104.084
11.3142 | -104.0677
11.3179 | -104.0552
11.3216 | -104.0446
11.3254 | -104.0391
11.3291 | -104.0317
11.3329 | -104.0296
11.3366 | -104.0394
11.3404 | -104.0412
11.3441 | -104.042
11.3478 | -104.0452
11.3516 | -104.0481
11.3553 | -104.0574
11.3591 | -104.0723
11.3628 | -104.0841
11.3666 | -104.0933
11.3703 | -104.1075
11.374 | -104.1238
11.3778 | -104.1375
11.3815 | -104.1533
11.3853 | -104.175
11.389 | -104.1984
11.3927 | -104.2196
11.3965 | -104.2408
11.4002 | -104.2633
11.404 | -104.2871
11.4077 | -104.3141
11.4115 | -104.3399
11.4152 | -104.3639
11.4189 | -104.3922
11.4227 | -104.4176
11.4264 | -104.4429
11.4302 | -104.466
11.4339 | -104.4849
11.4377 | -104.5037
11.4414 | -104.5236
11.4451 | -104.5433
11.4489 | -104.5595
11.4526 | -104.5781
11.4564 | -104.5904
11.4601 | -104.5965
11.4639 | -104.6039
11.4676 | -104.6162
11.4713 | -104.6261
11.4751 | -104.6351
11.4788 | -104.641
11.4826 | -104.6477
11.4863 | -104.6519
11.4901 | -104.6569
11.4938 | -104.6639
11.4975 | -104.669
11.5013 | -104.6734
11.505 | -104.6743
11.5088 | -104.6743
11.5125 | -104.6758
11.5163 | -104.676
11.52 | -104.6785
11.5237 | -104.6788
11.5275 | -104.673
11.5312 | -104.6682
11.535 | -104.666
11.5387 | -104.6602
11.5425 | -104.6552
11.5462 | -104.6561
11.5499 | -104.6564
11.5537 | -104.6533
11.5574 | -104.6541
11.5612 | -104.6557
11.5649 | -104.6579
11.5687 | -104.6611
11.5724 | -104.6651
11.5761 | -104.667
11.5799 | -104.6728
11.5836 | -104.6843
11.5874 | -104.6974
11.5911 | -104.7129
11.5949 | -104.7284
11.5986 | -104.7459
11.6023 | -104.764
11.6061 | -104.7802
11.6098 | -104.7967
11.6136 | -104.8139
11.6173 | -104.8331
11.6211 | -104.8484
11.6248 | -104.8655
11.6285 | -104.883
11.6323 | -104.891
11.636 | -104.8927
11.6398 | -104.8869
11.6435 | -104.8826
11.6473 | -104.8793
11.651 | -104.8658
11.6547 | -104.8452
11.6585 | -104.8239
11.6622 | -104.8015
11.666 | -104.7777
11.6697 | -104.746
11.6735 | -104.7145
11.6772 | -104.6755
11.6809 | -104.6315
11.6847 | -104.5868
11.6884 | -104.5399
11.6922 | -104.4913
11.6959 | -104.4406
11.6996 | -104.3947
11.7034 | -104.3498
11.7071 | -104.3061
11.7109 | -104.2642
11.7146 | -104.2222
11.7184 | -104.1873
11.7221 | -104.1552
11.7258 | -104.1226
11.7296 | -104.1002
11.7333 | -104.0876
11.7371 | -104.0769
11.7408 | -104.0682
11.7446 | -104.0627
11.7483 | -104.0604
11.752 | -104.0594
11.7558 | -104.0609
11.7595 | -104.0669
11.7633 | -104.0749
11.767 | -104.0808
11.7708 | -104.0904
11.7745 | -104.0997
11.7782 | -104.1077
11.782 | -104.1135
11.7857 | -104.1138
11.7895 | -104.1127
11.7932 | -104.1048
11.797 | -104.0973
11.8007 | -104.0927
11.8044 | -104.0863
11.8082 | -104.0785
11.8119 | -104.0648
11.8157 | -104.0484
11.8194 | -104.0373
11.8232 | -104.0307
11.8269 | -104.0198
11.8306 | -104.0072
11.8344 | -103.9925
11.8381 | -103.9812
11.8419 | -103.9709
11.8456 | -103.9679
11.8494 | -103.9662
11.8531 | -103.9617
11.8568 | -103.9601
11.8606 | -103.9612
11.8643 | -103.9627
11.8681 | -103.966
11.8718 | -103.974
11.8756 | -103.9943
11.8793 | -104.0069
11.883 | -104.0208
11.8868 | -104.0368
11.8905 | -104.0562
11.8943 | -104.0769
11.898 | -104.0995
11.9018 | -104.1214
11.9055 | -104.1371
11.9092 | -104.1511
11.913 | -104.1684
11.9167 | -104.1865
11.9205 | -104.2041
11.9242 | -104.2154
11.928 | -104.2239
11.9317 | -104.2303
11.9354 | -104.2365
11.9392 | -104.2447
11.9429 | -104.2504
11.9467 | -104.2565
11.9504 | -104.2623
11.9542 | -104.2684
11.9579 | -104.2772
11.9616 | -104.286
11.9654 | -104.2996
11.9691 | -104.3093
11.9729 | -104.3176
11.9766 | -104.3282
11.9804 | -104.3381
11.9841 | -104.3497
11.9878 | -104.3604
11.9916 | -104.3666
11.9953 | -104.3712
11.9991 | -104.3783
12.0028 | -104.3854
12.0066 | -104.3913
12.0103 | -104.3965
12.014 | -104.3983
12.0178 | -104.3944
12.0215 | -104.3921
12.0253 | -104.3912
12.029 | -104.3851
12.0327 | -104.3763
12.0365 | -104.3704
12.0402 | -104.3624
12.044 | -104.3532
12.0477 | -104.3464
12.0515 | -104.3392
12.0552 | -104.3342
12.0589 | -104.3345
12.0627 | -104.3337
12.0664 | -104.3253
12.0702 | -104.3139
12.0739 | -104.2995
12.0777 | -104.279
12.0814 | -104.2455
12.0851 | -104.2283
12.0889 | -104.2083
12.0926 | -104.1766
12.0964 | -104.1414
12.1001 | -104.1053
12.1039 | -104.0617
12.1076 | -104.0098
12.1113 | -103.9491
12.1151 | -103.8816
12.1188 | -103.8133
12.1226 | -103.7421
12.1263 | -103.6676
12.1301 | -103.5927
12.1338 | -103.5177
12.1375 | -103.4431
12.1413 | -103.3714
12.145 | -103.3074
12.1488 | -103.247
12.1525 | -103.1933
12.1563 | -103.1512
12.16 | -103.1173
12.1637 | -103.0907
12.1675 | -103.0739
12.1712 | -103.066
12.175 | -103.0657
12.1787 | -103.0806
12.1825 | -103.1069
12.1862 | -103.1398
12.1899 | -103.1782
12.1937 | -103.2288
12.1974 | -103.2844
12.2012 | -103.3442
12.2049 | -103.4042
12.2087 | -103.4603
12.2124 | -103.5148
12.2161 | -103.5679
12.2199 | -103.6184
12.2236 | -103.6641
12.2274 | -103.702
12.2311 | -103.7327
12.2349 | -103.7568
12.2386 | -103.7793
12.2423 | -103.7995
12.2461 | -103.8158
12.2498 | -103.8238
12.2536 | -103.8267
12.2573 | -103.8351
12.2611 | -103.8208
12.2648 | -103.8025
12.2685 | -103.7922
12.2723 | -103.7774
12.276 | -103.7634
12.2798 | -103.7528
12.2835 | -103.7436
12.2873 | -103.7336
12.291 | -103.7264
12.2947 | -103.7198
12.2985 | -103.715
12.3022 | -103.711
12.306 | -103.7075
12.3097 | -103.704
12.3134 | -103.7047
12.3172 | -103.7059
12.3209 | -103.7063
12.3247 | -103.71
12.3284 | -103.7162
12.3322 | -103.7238
12.3359 | -103.7316
12.3396 | -103.7438
12.3434 | -103.7573
12.3471 | -103.7735
12.3509 | -103.7949
12.3546 | -103.8114
12.3584 | -103.8278
12.3621 | -103.847
12.3658 | -103.8696
12.3696 | -103.8909
12.3733 | -103.916
12.3771 | -103.9422
12.3808 | -103.9704
12.3846 | -104.0009
12.3883 | -104.0298
12.392 | -104.0549
12.3958 | -104.0767
12.3995 | -104.0987
12.4033 | -104.1193
12.407 | -104.1354
12.4108 | -104.149
12.4145 | -104.1589
12.4182 | -104.1552
12.422 | -104.1616
12.4257 | -104.1686
12.4295 | -104.1757
12.4332 | -104.179
12.437 | -104.1789
12.4407 | -104.1785
12.4444 | -104.1751
12.4482 | -104.1692
12.4519 | -104.1642
12.4557 | -104.1568
12.4594 | -104.1512
12.4632 | -104.1491
12.4669 | -104.1497
12.4706 | -104.1438
12.4744 | -104.1362
12.4781 | -104.1319
12.4819 | -104.1319
12.4856 | -104.1322
12.4894 | -104.1364
12.4931 | -104.1387
12.4968 | -104.1392
12.5006 | -104.1421
12.5043 | -104.1468
12.5081 | -104.1489
12.5118 | -104.1536
12.5156 | -104.1592
12.5193 | -104.1609
12.523 | -104.1627
12.5268 | -104.1635
12.5305 | -104.1596
12.5343 | -104.1537
12.538 | -104.1512
12.5418 | -104.1471
12.5455 | -104.14
12.5492 | -104.1313
12.553 | -104.1193
12.5567 | -104.1101
12.5605 | -104.1053
12.5642 | -104.0986
12.568 | -104.09
12.5717 | -104.0836
12.5754 | -104.0821
12.5792 | -104.0812
12.5829 | -104.084
12.5867 | -104.0925
12.5904 | -104.1009
12.5942 | -104.1137
12.5979 | -104.1288
12.6016 | -104.1444
12.6054 | -104.1624
12.6091 | -104.1812
12.6129 | -104.204
12.6166 | -104.2272
12.6204 | -104.2525
12.6241 | -104.2815
12.6278 | -104.312
12.6316 | -104.3402
12.6353 | -104.3675
12.6391 | -104.3934
12.6428 | -104.4195
12.6465 | -104.4392
12.6503 | -104.4574
12.654 | -104.4773
12.6578 | -104.4923
12.6615 | -104.4993
12.6653 | -104.5036
12.669 | -104.5078
12.6727 | -104.5051
12.6765 | -104.4956
12.6802 | -104.4868
12.684 | -104.4722
12.6877 | -104.4534
12.6915 | -104.4317
12.6952 | -104.4037
12.6989 | -104.3772
12.7027 | -104.3517
12.7064 | -104.325
12.7102 | -104.2962
12.7139 | -104.2658
12.7177 | -104.237
12.7214 | -104.2076
12.7251 | -104.1801
12.7289 | -104.1569
12.7326 | -104.1371
12.7364 | -104.1202
12.7401 | -104.1033
12.7439 | -104.0879
12.7476 | -104.0794
12.7513 | -104.072
12.7551 | -104.0658
12.7588 | -104.0625
12.7626 | -104.0626
12.7663 | -104.062
12.7701 | -104.0623
12.7738 | -104.0693
12.7775 | -104.0783
12.7813 | -104.088
12.785 | -104.1031
12.7888 | -104.1157
12.7925 | -104.1276
12.7963 | -104.1386
12.8 | -104.1485
12.8037 | -104.156
12.8075 | -104.1621
12.8112 | -104.1677
12.815 | -104.1701
12.8187 | -104.1701
12.8225 | -104.1697
12.8262 | -104.1649
12.8299 | -104.1618
12.8337 | -104.1558
12.8374 | -104.1486
12.8412 | -104.144
12.8449 | -104.1428
12.8487 | -104.1391
12.8524 | -104.1356
12.8561 | -104.1345
12.8599 | -104.14
12.8636 | -104.1465
12.8674 | -104.1571
12.8711 | -104.1708
12.8749 | -104.1886
12.8786 | -104.2078
12.8823 | -104.2287
12.8861 | -104.2482
12.8898 | -104.2666
12.8936 | -104.2901
12.8973 | -104.3187
12.9011 | -104.3468
12.9048 | -104.3714
12.9085 | -104.396
12.9123 | -104.4172
12.916 | -104.4359
12.9198 | -104.453
12.9235 | -104.4719
12.9273 | -104.49
12.931 | -104.5044
12.9347 | -104.5174
12.9385 | -104.5246
12.9422 | -104.5298
12.946 | -104.5375
12.9497 | -104.544
12.9535 | -104.5453
12.9572 | -104.5318
12.9609 | -104.5258
12.9647 | -104.519
12.9684 | -104.5133
12.9722 | -104.5018
12.9759 | -104.4934
12.9796 | -104.486
12.9834 | -104.4736
12.9871 | -104.4587
12.9909 | -104.4436
12.9946 | -104.4298
12.9984 | -104.4144
13.0021 | -104.397
13.0058 | -104.3778
13.0096 | -104.3572
13.0133 | -104.3393
13.0171 | -104.3222
13.0208 | -104.3014
13.0246 | -104.2776
13.0283 | -104.2547
13.032 | -104.2307
13.0358 | -104.2111
13.0395 | -104.188
13.0433 | -104.165
13.047 | -104.1399
13.0508 | -104.1129
13.0545 | -104.0841
13.0582 | -104.0527
13.062 | -104.0261
13.0657 | -104.0016
13.0695 | -103.9745
13.0732 | -103.9516
13.077 | -103.9311
13.0807 | -103.9074
13.0844 | -103.882
13.0882 | -103.8615
13.0919 | -103.8447
13.0957 | -103.831
13.0994 | -103.8189
13.1032 | -103.8069
13.1069 | -103.7966
13.1106 | -103.7886
13.1144 | -103.7847
13.1181 | -103.782
13.1219 | -103.7833
13.1256 | -103.7877
13.1294 | -103.7947
13.1331 | -103.802
13.1368 | -103.8126
13.1406 | -103.8255
13.1443 | -103.8424
13.1481 | -103.859
13.1518 | -103.879
13.1556 | -103.8986
13.1593 | -103.921
13.163 | -103.9482
13.1668 | -103.9752
13.1705 | -104.002
13.1743 | -104.0347
13.178 | -104.07
13.1818 | -104.1027
13.1855 | -104.1293
13.1892 | -104.1604
13.193 | -104.1955
13.1967 | -104.2301
13.2005 | -104.2616
13.2042 | -104.2871
13.208 | -104.3141
13.2117 | -104.3449
13.2154 | -104.3735
13.2192 | -104.3958
13.2229 | -104.4173
13.2267 | -104.4373
13.2304 | -104.4574
13.2342 | -104.4753
13.2379 | -104.4926
13.2416 | -104.5094
13.2454 | -104.5257
13.2491 | -104.5403
13.2529 | -104.5499
13.2566 | -104.5536
13.2604 | -104.5518
13.2641 | -104.5465
13.2678 | -104.5407
13.2716 | -104.5358
13.2753 | -104.5284
13.2791 | -104.5173
13.2828 | -104.4983
13.2866 | -104.4775
13.2903 | -104.4558
13.294 | -104.4305
13.2978 | -104.4048
13.3015 | -104.3755
13.3053 | -104.3413
13.309 | -104.3085
13.3127 | -104.2736
13.3165 | -104.2384
13.3202 | -104.1992
13.324 | -104.161
13.3277 | -104.1218
13.3315 | -104.0838
13.3352 | -104.0512
13.3389 | -104.0205
13.3427 | -103.9901
13.3464 | -103.9642
13.3502 | -103.9396
13.3539 | -103.9164
13.3577 | -103.898
13.3614 | -103.8849
13.3651 | -103.8754
13.3689 | -103.8722
13.3726 | -103.8746
13.3764 | -103.8825
13.3801 | -103.8896
13.3839 | -103.8987
13.3876 | -103.9103
13.3913 | -103.9248
13.3951 | -103.9405
13.3988 | -103.9563
13.4026 | -103.9731
13.4063 | -103.9931
13.4101 | -104.0144
13.4138 | -104.0352
13.4175 | -104.0552
13.4213 | -104.0767
13.425 | -104.0978
13.4288 | -104.1175
13.4325 | -104.1401
13.4363 | -104.1622
13.44 | -104.1846
13.4437 | -104.2049
13.4475 | -104.2278
13.4512 | -104.251
13.455 | -104.2716
13.4587 | -104.2894
13.4625 | -104.3048
13.4662 | -104.318
13.4699 | -104.3288
13.4737 | -104.3393
13.4774 | -104.3437
13.4812 | -104.3452
13.4849 | -104.3457
13.4887 | -104.3424
13.4924 | -104.3325
13.4961 | -104.3079
13.4999 | -104.295
13.5036 | -104.2757
13.5074 | -104.2519
13.5111 | -104.2224
13.5149 | -104.1922
13.5186 | -104.16
13.5223 | -104.1311
13.5261 | -104.0988
13.5298 | -104.0619
13.5336 | -104.0268
13.5373 | -103.9922
13.5411 | -103.9559
13.5448 | -103.9219
13.5485 | -103.8903
13.5523 | -103.8591
13.556 | -103.8302
13.5598 | -103.8004
13.5635 | -103.7719
13.5673 | -103.7492
13.571 | -103.7308
13.5747 | -103.7157
13.5785 | -103.7032
13.5822 | -103.6956
13.586 | -103.6905
13.5897 | -103.6881
13.5935 | -103.685
13.5972 | -103.6857
13.6009 | -103.6863
13.6047 | -103.6919
13.6084 | -103.6997
13.6122 | -103.7076
13.6159 | -103.7179
13.6196 | -103.7265
13.6234 | -103.7372
13.6271 | -103.7452
13.6309 | -103.7556
13.6346 | -103.768
13.6384 | -103.7789
13.6421 | -103.7909
13.6458 | -103.8071
13.6496 | -103.8199
13.6533 | -103.8314
13.6571 | -103.8447
13.6608 | -103.8574
13.6646 | -103.8674
13.6683 | -103.8783
13.672 | -103.8873
13.6758 | -103.8981
13.6795 | -103.9085
13.6833 | -103.9173
13.687 | -103.9293
13.6908 | -103.9402
13.6945 | -103.9479
13.6982 | -103.9569
13.702 | -103.9681
13.7057 | -103.9809
13.7095 | -103.9964
13.7132 | -104.01
13.717 | -104.0218
13.7207 | -104.0307
13.7244 | -104.0365
13.7282 | -104.0418
13.7319 | -104.0479
13.7357 | -104.0535
13.7394 | -104.0588
13.7432 | -104.0636
13.7469 | -104.0688
13.7506 | -104.0732
13.7544 | -104.0759
13.7581 | -104.0799
13.7619 | -104.0822
13.7656 | -104.0792
13.7694 | -104.0738
13.7731 | -104.0701
13.7768 | -104.0659
13.7806 | -104.0605
13.7843 | -104.0545
13.7881 | -104.0457
13.7918 | -104.0364
13.7956 | -104.0249
13.7993 | -104.0157
13.803 | -104.0036
13.8068 | -103.9924
13.8105 | -103.9817
13.8143 | -103.9705
13.818 | -103.9645
13.8218 | -103.9563
13.8255 | -103.95
13.8292 | -103.9436
13.833 | -103.9363
13.8367 | -103.9293
13.8405 | -103.9196
13.8442 | -103.9058
13.848 | -103.9019
13.8517 | -103.9
13.8554 | -103.8979
13.8592 | -103.8952
13.8629 | -103.8961
13.8667 | -103.8971
13.8704 | -103.8994
13.8742 | -103.8955
13.8779 | -103.8922
13.8816 | -103.8889
13.8854 | -103.8801
13.8891 | -103.8727
13.8929 | -103.8637
13.8966 | -103.856
13.9004 | -103.849
13.9041 | -103.8348
13.9078 | -103.813
13.9116 | -103.7928
13.9153 | -103.7761
13.9191 | -103.76
13.9228 | -103.7483
13.9266 | -103.7358
13.9303 | -103.7191
13.934 | -103.7027
13.9378 | -103.6885
13.9415 | -103.6726
13.9453 | -103.6535
13.949 | -103.6351
13.9527 | -103.6167
13.9565 | -103.5981
13.9602 | -103.5781
13.964 | -103.5464
13.9677 | -103.5151
13.9715 | -103.4929
13.9752 | -103.4704
13.9789 | -103.4444
13.9827 | -103.4197
13.9864 | -103.3968
13.9902 | -103.3721
13.9939 | -103.3468
13.9977 | -103.3204
14.0014 | -103.2944
14.0051 | -103.2683
14.0089 | -103.2417
14.0126 | -103.2179
14.0164 | -103.1963
14.0201 | -103.1768
14.0239 | -103.1571
14.0276 | -103.1365
14.0313 | -103.1196
14.0351 | -103.1217
14.0388 | -103.119
14.0426 | -103.1198
14.0463 | -103.1179
14.0501 | -103.112
14.0538 | -103.1125
14.0575 | -103.1182
14.0613 | -103.1311
14.065 | -103.1453
14.0688 | -103.1603
14.0725 | -103.1787
14.0763 | -103.202
14.08 | -103.2254
14.0837 | -103.249
14.0875 | -103.2742
14.0912 | -103.2963
14.095 | -103.3203
14.0987 | -103.344
14.1025 | -103.3671
14.1062 | -103.3945
14.1099 | -103.4225
14.1137 | -103.444
14.1174 | -103.4609
14.1212 | -103.4759
14.1249 | -103.4879
14.1287 | -103.5028
14.1324 | -103.5211
14.1361 | -103.5424
14.1399 | -103.5629
14.1436 | -103.5858
14.1474 | -103.6104
14.1511 | -103.6375
14.1549 | -103.6657
14.1586 | -103.6958
14.1623 | -103.729
14.1661 | -103.7756
14.1698 | -103.823
14.1736 | -103.8582
14.1773 | -103.889
14.1811 | -103.9194
14.1848 | -103.9502
14.1885 | -103.9848
14.1923 | -104.0214
14.196 | -104.0538
14.1998 | -104.0828
14.2035 | -104.1082
14.2073 | -104.1239
14.211 | -104.1362
14.2147 | -104.1465
14.2185 | -104.1522
14.2222 | -104.1518
14.226 | -104.1462
14.2297 | -104.1335
14.2334 | -104.1124
14.2372 | -104.0871
14.2409 | -104.0561
14.2447 | -104.0165
14.2484 | -103.9744
14.2522 | -103.9388
14.2559 | -103.9003
14.2596 | -103.8525
14.2634 | -103.7999
14.2671 | -103.7513
14.2709 | -103.7048
14.2746 | -103.6616
14.2784 | -103.6199
14.2821 | -103.5803
14.2858 | -103.5459
14.2896 | -103.5189
14.2933 | -103.5007
14.2971 | -103.4853
14.3008 | -103.4712
14.3046 | -103.4585
14.3083 | -103.45
14.312 | -103.4478
14.3158 | -103.4572
14.3195 | -103.4707
14.3233 | -103.4845
14.327 | -103.499
14.3308 | -103.5192
14.3345 | -103.5372
14.3382 | -103.5581
14.342 | -103.581
14.3457 | -103.607
14.3495 | -103.6316
14.3532 | -103.653
14.357 | -103.6718
14.3607 | -103.6902
14.3644 | -103.7074
14.3682 | -103.7234
14.3719 | -103.7396
14.3757 | -103.752
14.3794 | -103.7643
14.3832 | -103.7798
14.3869 | -103.7946
14.3906 | -103.8041
14.3944 | -103.8081
14.3981 | -103.8101
14.4019 | -103.8122
14.4056 | -103.8146
14.4094 | -103.8169
14.4131 | -103.8236
14.4168 | -103.8311
14.4206 | -103.8335
14.4243 | -103.8321
14.4281 | -103.8317
14.4318 | -103.8358
14.4356 | -103.8463
14.4393 | -103.8596
14.443 | -103.8723
14.4468 | -103.8833
14.4505 | -103.8976
14.4543 | -103.912
14.458 | -103.9265
14.4618 | -103.9437
14.4655 | -103.964
14.4692 | -103.9883
14.473 | -104.0078
14.4767 | -104.0225
14.4805 | -104.0378
14.4842 | -104.0546
14.488 | -104.0755
14.4917 | -104.0908
14.4954 | -104.0998
14.4992 | -104.1063
14.5029 | -104.114
14.5067 | -104.1203
14.5104 | -104.1274
14.5142 | -104.1343
14.5179 | -104.1349
14.5216 | -104.1285
14.5254 | -104.1228
14.5291 | -104.116
14.5329 | -104.1118
14.5366 | -104.1043
14.5404 | -104.0931
14.5441 | -104.0787
14.5478 | -104.0641
14.5516 | -104.0491
14.5553 | -104.0324
14.5591 | -104.0184
14.5628 | -104.0083
14.5665 | -103.9979
14.5703 | -103.9826
14.574 | -103.9653
14.5778 | -103.9568
14.5815 | -103.9415
14.5853 | -103.927
14.589 | -103.9176
14.5927 | -103.9069
14.5965 | -103.8994
14.6002 | -103.8942
14.604 | -103.8869
14.6077 | -103.8808
14.6115 | -103.8778
14.6152 | -103.8774
14.6189 | -103.8764
14.6227 | -103.8752
14.6264 | -103.8736
14.6302 | -103.8742
14.6339 | -103.8771
14.6377 | -103.8805
14.6414 | -103.8801
14.6451 | -103.8778
14.6489 | -103.8765
14.6526 | -103.8742
14.6564 | -103.8743
14.6601 | -103.8769
14.6639 | -103.8815
14.6676 | -103.8853
14.6713 | -103.889
14.6751 | -103.8938
14.6788 | -103.8978
14.6826 | -103.9033
14.6863 | -103.9128
14.6901 | -103.9218
14.6938 | -103.9293
14.6975 | -103.9337
14.7013 | -103.9417
14.705 | -103.9512
14.7088 | -103.9596
14.7125 | -103.969
14.7163 | -103.9822
14.72 | -103.9952
14.7237 | -104.0049
14.7275 | -104.0181
14.7312 | -104.0313
14.735 | -104.0463
14.7387 | -104.0625
14.7425 | -104.0766
14.7462 | -104.0899
14.7499 | -104.1058
14.7537 | -104.1201
14.7574 | -104.1328
14.7612 | -104.145
14.7649 | -104.1617
14.7687 | -104.178
14.7724 | -104.19
14.7761 | -104.202
14.7799 | -104.2172
14.7836 | -104.2328
14.7874 | -104.244
14.7911 | -104.2532
14.7949 | -104.2652
14.7986 | -104.2732
14.8023 | -104.2785
14.8061 | -104.2852
14.8098 | -104.2913
14.8136 | -104.2961
14.8173 | -104.3022
14.8211 | -104.3028
14.8248 | -104.3026
14.8285 | -104.302
14.8323 | -104.303
14.836 | -104.3053
14.8398 | -104.3068
14.8435 | -104.3047
14.8473 | -104.3028
14.851 | -104.3023
14.8547 | -104.3028
14.8585 | -104.303
14.8622 | -104.2981
14.866 | -104.2961
14.8697 | -104.2962
14.8735 | -104.297
14.8772 | -104.296
14.8809 | -104.2943
14.8847 | -104.2948
14.8884 | -104.2992
14.8922 | -104.306
14.8959 | -104.3118
14.8996 | -104.3185
14.9034 | -104.3253
14.9071 | -104.3322
14.9109 | -104.3408
14.9146 | -104.3525
14.9184 | -104.3639
14.9221 | -104.377
14.9258 | -104.3881
14.9296 | -104.3971
14.9333 | -104.4073
14.9371 | -104.4197
14.9408 | -104.4312
14.9446 | -104.4435
14.9483 | -104.4543
14.952 | -104.4661
14.9558 | -104.478
14.9595 | -104.4892
14.9633 | -104.5
14.967 | -104.5116
14.9708 | -104.5222
14.9745 | -104.5289
14.9782 | -104.5345
14.982 | -104.5422
14.9857 | -104.5438
14.9895 | -104.5085
14.9932 | -104.4772
14.997 | -104.4631
15.0007 | -104.4476
15.0044 | -104.427
15.0082 | -104.4017
15.0119 | -104.3766
15.0157 | -104.3443
15.0194 | -104.3087
15.0232 | -104.2663
15.0269 | -104.2219
15.0306 | -104.1735
15.0344 | -104.1215
15.0381 | -104.0658
15.0419 | -104.0077
15.0456 | -103.9499
15.0494 | -103.8886
15.0531 | -103.8277
15.0568 | -103.7695
15.0606 | -103.7107
15.0643 | -103.6556
15.0681 | -103.6055
15.0718 | -103.559
15.0756 | -103.5177
15.0793 | -103.4773
15.083 | -103.4389
15.0868 | -103.4084
15.0905 | -103.3869
15.0943 | -103.3719
15.098 | -103.3616
15.1018 | -103.3551
15.1055 | -103.3555
15.1092 | -103.3644
15.113 | -103.3796
15.1167 | -103.4174
15.1205 | -103.4377
15.1242 | -103.4598
15.128 | -103.4832
15.1317 | -103.5088
15.1354 | -103.5394
15.1392 | -103.5702
15.1429 | -103.601
15.1467 | -103.6311
15.1504 | -103.6607
15.1542 | -103.6913
15.1579 | -103.7195
15.1616 | -103.7478
15.1654 | -103.7748
15.1691 | -103.8003
15.1729 | -103.8229
15.1766 | -103.845
15.1804 | -103.8658
15.1841 | -103.885
15.1878 | -103.9046
15.1916 | -103.9229
15.1953 | -103.9382
15.1991 | -103.9508
15.2028 | -103.967
15.2066 | -104.0083
15.2103 | -104.0352
15.214 | -104.0401
15.2178 | -104.0443
15.2215 | -104.0474
15.2253 | -104.0504
15.229 | -104.0553
15.2327 | -104.059
15.2365 | -104.0665
15.2402 | -104.0744
15.244 | -104.0796
15.2477 | -104.0871
15.2515 | -104.097
15.2552 | -104.1037
15.2589 | -104.1043
15.2627 | -104.1019
15.2664 | -104.1046
15.2702 | -104.1061
15.2739 | -104.1047
15.2777 | -104.1015
15.2814 | -104.096
15.2851 | -104.0873
15.2889 | -104.0771
15.2926 | -104.0657
15.2964 | -104.0548
15.3001 | -104.0442
15.3039 | -104.0303
15.3076 | -104.0136
15.3113 | -103.9968
15.3151 | -103.9776
15.3188 | -103.9582
15.3226 | -103.9399
15.3263 | -103.923
15.3301 | -103.9054
15.3338 | -103.8904
15.3375 | -103.8741
15.3413 | -103.8561
15.345 | -103.8427
15.3488 | -103.834
15.3525 | -103.8224
15.3563 | -103.8087
15.36 | -103.7975
15.3637 | -103.7881
15.3675 | -103.7819
15.3712 | -103.7771
15.375 | -103.7706
15.3787 | -103.7646
15.3825 | -103.7588
15.3862 | -103.7543
15.3899 | -103.7533
15.3937 | -103.7502
15.3974 | -103.7442
15.4012 | -103.7348
15.4049 | -103.7252
15.4087 | -103.7145
15.4124 | -103.7095
15.4161 | -103.702
15.4199 | -103.6944
15.4236 | -103.6877
15.4274 | -103.6784
15.4311 | -103.6641
15.4349 | -103.6546
15.4386 | -103.6462
15.4423 | -103.6379
15.4461 | -103.6276
15.4498 | -103.6143
15.4536 | -103.6049
15.4573 | -103.5913
15.4611 | -103.5772
15.4648 | -103.5662
15.4685 | -103.5553
15.4723 | -103.5444
15.476 | -103.531
15.4798 | -103.5202
15.4835 | -103.5132
15.4873 | -103.5056
15.491 | -103.4987
15.4947 | -103.4927
15.4985 | -103.4864
15.5022 | -103.4804
15.506 | -103.4754
15.5097 | -103.4692
15.5135 | -103.4688
15.5172 | -103.4683
15.5209 | -103.4692
15.5247 | -103.4711
15.5284 | -103.4724
15.5322 | -103.4735
15.5359 | -103.4772
15.5396 | -103.4817
15.5434 | -103.4837
15.5471 | -103.4832
15.5509 | -103.4852
15.5546 | -103.4866
15.5584 | -103.4904
15.5621 | -103.4944
15.5658 | -103.4905
15.5696 | -103.4823
15.5733 | -103.4798
15.5771 | -103.4776
15.5808 | -103.4746
15.5846 | -103.4722
15.5883 | -103.4633
15.592 | -103.4524
15.5958 | -103.4414
15.5995 | -103.4364
15.6033 | -103.4287
15.607 | -103.421
15.6108 | -103.4137
15.6145 | -103.4085
15.6182 | -103.4041
15.622 | -103.402
15.6257 | -103.3997
15.6295 | -103.3966
15.6332 | -103.3857
15.637 | -103.3745
15.6407 | -103.3687
15.6444 | -103.3619
15.6482 | -103.3583
15.6519 | -103.3563
15.6557 | -103.3477
15.6594 | -103.3423
15.6632 | -103.3403
15.6669 | -103.3404
15.6706 | -103.341
15.6744 | -103.3412
15.6781 | -103.3416
15.6819 | -103.3452
15.6856 | -103.352
15.6894 | -103.3584
15.6931 | -103.3625
15.6968 | -103.3685
15.7006 | -103.3829
15.7043 | -103.3978
15.7081 | -103.4055
15.7118 | -103.4146
15.7156 | -103.4292
15.7193 | -103.4438
15.723 | -103.46
15.7268 | -103.4748
15.7305 | -103.4902
15.7343 | -103.5088
15.738 | -103.5211
15.7418 | -103.5339
15.7455 | -103.5491
15.7492 | -103.5626
15.753 | -103.5799
15.7567 | -103.5968
15.7605 | -103.6132
15.7642 | -103.6309
15.768 | -103.6496
15.7717 | -103.6689
15.7754 | -103.6836
15.7792 | -103.7004
15.7829 | -103.7168
15.7867 | -103.7333
15.7904 | -103.7541
15.7942 | -103.7751
15.7979 | -103.7925
15.8016 | -103.8079
15.8054 | -103.822
15.8091 | -103.8376
15.8129 | -103.8581
15.8166 | -103.8746
15.8204 | -103.8888
15.8241 | -103.9005
15.8278 | -103.9156
15.8316 | -103.9304
15.8353 | -103.9416
15.8391 | -103.9525
15.8428 | -103.9663
15.8466 | -103.9817
15.8503 | -103.9959
15.854 | -104.0064
15.8578 | -104.0185
15.8615 | -104.032
15.8653 | -104.038
15.869 | -104.0446
15.8727 | -104.0497
15.8765 | -104.0572
15.8802 | -104.0657
15.884 | -104.0678
15.8877 | -104.0704
15.8915 | -104.0664
15.8952 | -104.0618
15.8989 | -104.0574
15.9027 | -104.0513
15.9064 | -104.0408
15.9102 | -104.0274
15.9139 | -104.0134
15.9177 | -104.0043
15.9214 | -103.9982
15.9251 | -103.9906
15.9289 | -103.9795
15.9326 | -103.9679
15.9364 | -103.9585
15.9401 | -103.944
15.9439 | -103.9252
15.9476 | -103.909
15.9513 | -103.8949
15.9551 | -103.8784
15.9588 | -103.8617
15.9626 | -103.8446
15.9663 | -103.8288
15.9701 | -103.811
15.9738 | -103.792
15.9775 | -103.777
15.9813 | -103.7621
15.985 | -103.7479
15.9888 | -103.7321
15.9925 | -103.7166
15.9963 | -103.7035
16 | -103.6903
16.0037 | -103.6815
16.0075 | -103.6707
16.0112 | -103.6616
16.015 | -103.653
16.0187 | -103.645
16.0225 | -103.6399
16.0262 | -103.6356
16.0299 | -103.6365
16.0337 | -103.6367
16.0374 | -103.6374
16.0412 | -103.6409
16.0449 | -103.6479
16.0487 | -103.6567
16.0524 | -103.6627
16.0561 | -103.6701
16.0599 | -103.6781
16.0636 | -103.6888
16.0674 | -103.7033
16.0711 | -103.7149
16.0749 | -103.7245
16.0786 | -103.7359
16.0823 | -103.7498
16.0861 | -103.7629
16.0898 | -103.7759
16.0936 | -103.7858
16.0973 | -103.7912
16.1011 | -103.7917
16.1048 | -103.7906
16.1085 | -103.7902
16.1123 | -103.7883
16.116 | -103.7826
16.1198 | -103.7782
16.1235 | -103.7718
16.1273 | -103.7631
16.131 | -103.7539
16.1347 | -103.7427
16.1385 | -103.7306
16.1422 | -103.7148
16.146 | -103.6931
16.1497 | -103.6727
16.1534 | -103.6517
16.1572 | -103.6302
16.1609 | -103.6123
16.1647 | -103.5961
16.1684 | -103.581
16.1722 | -103.564
16.1759 | -103.5456
16.1796 | -103.5296
16.1834 | -103.5156
16.1871 | -103.5036
16.1909 | -103.4926
16.1946 | -103.4819
16.1984 | -103.4828
16.2021 | -103.4731
16.2058 | -103.464
16.2096 | -103.4576
16.2133 | -103.4554
16.2171 | -103.4586
16.2208 | -103.4643
16.2246 | -103.4698
16.2283 | -103.4763
16.232 | -103.4882
16.2358 | -103.4986
16.2395 | -103.5098
16.2433 | -103.5214
16.247 | -103.5321
16.2508 | -103.5438
16.2545 | -103.5567
16.2582 | -103.5693
16.262 | -103.5761
16.2657 | -103.5808
16.2695 | -103.5835
16.2732 | -103.5844
16.277 | -103.5864
16.2807 | -103.5896
16.2844 | -103.5915
16.2882 | -103.5883
16.2919 | -103.5842
16.2957 | -103.5812
16.2994 | -103.5759
16.3032 | -103.5686
16.3069 | -103.5621
16.3106 | -103.5534
16.3144 | -103.5429
16.3181 | -103.536
16.3219 | -103.5278
16.3256 | -103.5196
16.3294 | -103.5113
16.3331 | -103.5064
16.3368 | -103.5007
16.3406 | -103.4959
16.3443 | -103.4929
16.3481 | -103.4912
16.3518 | -103.49
16.3556 | -103.4897
16.3593 | -103.4927
16.363 | -103.4998
16.3668 | -103.5073
16.3705 | -103.5146
16.3743 | -103.5263
16.378 | -103.538
16.3818 | -103.5489
16.3855 | -103.563
16.3892 | -103.578
16.393 | -103.5915
16.3967 | -103.6055
16.4005 | -103.6206
16.4042 | -103.6404
16.408 | -103.6628
16.4117 | -103.6819
16.4154 | -103.7015
16.4192 | -103.7207
16.4229 | -103.7405
16.4267 | -103.7625
16.4304 | -103.783
16.4342 | -103.8041
16.4379 | -103.8253
16.4416 | -103.845
16.4454 | -103.8614
16.4491 | -103.8755
16.4529 | -103.885
16.4566 | -103.8934
16.4604 | -103.9025
16.4641 | -103.9088
16.4678 | -103.9135
16.4716 | -103.9208
16.4753 | -103.9271
16.4791 | -103.9335
16.4828 | -103.9373
16.4866 | -103.9392
16.4903 | -103.9377
16.494 | -103.9329
16.4978 | -103.9314
16.5015 | -103.9356
16.5053 | -103.9373
16.509 | -103.9355
16.5127 | -103.9318
16.5165 | -103.9231
16.5202 | -103.9158
16.524 | -103.9103
16.5277 | -103.9078
16.5315 | -103.9032
16.5352 | -103.9007
16.5389 | -103.9016
16.5427 | -103.9028
16.5464 | -103.904
16.5502 | -103.9069
16.5539 | -103.91
16.5577 | -103.914
16.5614 | -103.9209
16.5651 | -103.9306
16.5689 | -103.9412
16.5726 | -103.954
16.5764 | -103.9637
16.5801 | -103.9711
16.5839 | -103.9808
16.5876 | -103.9948
16.5913 | -104.0097
16.5951 | -104.0219
16.5988 | -104.0291
16.6026 | -104.0382
16.6063 | -104.0478
16.6101 | -104.057
16.6138 | -104.07
16.6175 | -104.0833
16.6213 | -104.0941
16.625 | -104.1062
16.6288 | -104.1187
16.6325 | -104.1301
16.6363 | -104.1409
16.64 | -104.1481
16.6437 | -104.1543
16.6475 | -104.1565
16.6512 | -104.1569
16.655 | -104.1568
16.6587 | -104.1564
16.6625 | -104.1579
16.6662 | -104.1582
16.6699 | -104.157
16.6737 | -104.1559
16.6774 | -104.1535
16.6812 | -104.15
16.6849 | -104.1455
16.6887 | -104.1434
16.6924 | -104.1422
16.6961 | -104.141
16.6999 | -104.1402
16.7036 | -104.1355
16.7074 | -104.1322
16.7111 | -104.1302
16.7149 | -104.1273
16.7186 | -104.1236
16.7223 | -104.1189
16.7261 | -104.1161
16.7298 | -104.1189
16.7336 | -104.1198
16.7373 | -104.1127
16.7411 | -104.1022
16.7448 | -104.0939
16.7485 | -104.0873
16.7523 | -104.0799
16.756 | -104.0746
16.7598 | -104.0667
16.7635 | -104.0591
16.7673 | -104.0506
16.771 | -104.0414
16.7747 | -104.0295
16.7785 | -104.0176
16.7822 | -104.0039
16.786 | -103.9888
16.7897 | -103.973
16.7934 | -103.9549
16.7972 | -103.9367
16.8009 | -103.919
16.8047 | -103.9023
16.8084 | -103.8824
16.8122 | -103.8605
16.8159 | -103.8403
16.8196 | -103.8205
16.8234 | -103.8012
16.8271 | -103.7822
16.8309 | -103.7613
16.8346 | -103.7456
16.8384 | -103.7354
16.8421 | -103.7239
16.8458 | -103.7132
16.8496 | -103.6999
16.8533 | -103.6842
16.8571 | -103.6711
16.8608 | -103.6621
16.8646 | -103.6532
16.8683 | -103.6431
16.872 | -103.6339
16.8758 | -103.6247
16.8795 | -103.6195
16.8833 | -103.6167
16.887 | -103.6154
16.8908 | -103.6092
16.8945 | -103.6047
16.8982 | -103.6027
16.902 | -103.5984
16.9057 | -103.5934
16.9095 | -103.5847
16.9132 | -103.5747
16.917 | -103.5681
16.9207 | -103.5621
16.9244 | -103.5575
16.9282 | -103.5482
16.9319 | -103.54
16.9357 | -103.5323
16.9394 | -103.5252
16.9432 | -103.5171
16.9469 | -103.505
16.9506 | -103.4945
16.9544 | -103.4834
16.9581 | -103.4751
16.9619 | -103.4648
16.9656 | -103.4512
16.9694 | -103.4342
16.9731 | -103.4175
16.9768 | -103.4021
16.9806 | -103.3906
16.9843 | -103.3796
16.9881 | -103.3672
16.9918 | -103.354
16.9956 | -103.3426
16.9993 | -103.3322
17.003 | -103.3225
17.0068 | -103.313
17.0105 | -103.3021
17.0143 | -103.2925
17.018 | -103.2856
17.0218 | -103.282
17.0255 | -103.2821
17.0292 | -103.2827
17.033 | -103.2845
17.0367 | -103.289
17.0405 | -103.2955
17.0442 | -103.3043
17.048 | -103.3117
17.0517 | -103.321
17.0554 | -103.3316
17.0592 | -103.3468
17.0629 | -103.3624
17.0667 | -103.3752
17.0704 | -103.3861
17.0742 | -103.398
17.0779 | -103.4063
17.0816 | -103.4124
17.0854 | -103.4173
17.0891 | -103.423
17.0929 | -103.4273
17.0966 | -103.4328
17.1004 | -103.4356
17.1041 | -103.4365
17.1078 | -103.4369
17.1116 | -103.4382
17.1153 | -103.4375
17.1191 | -103.4358
17.1228 | -103.4314
17.1266 | -103.4265
17.1303 | -103.4233
17.134 | -103.4208
17.1378 | -103.4123
17.1415 | -103.4048
17.1453 | -103.3997
17.149 | -103.3969
17.1527 | -103.3955
17.1565 | -103.3912
17.1602 | -103.384
17.164 | -103.3789
17.1677 | -103.3732
17.1715 | -103.372
17.1752 | -103.3695
17.1789 | -103.3669
17.1827 | -103.3658
17.1864 | -103.3647
17.1902 | -103.3654
17.1939 | -103.3659
17.1977 | -103.3673
17.2014 | -103.3719
17.2051 | -103.3775
17.2089 | -103.3812
17.2126 | -103.3843
17.2164 | -103.3928
17.2201 | -103.4025
17.2239 | -103.4086
17.2276 | -103.4118
17.2313 | -103.4168
17.2351 | -103.4258
17.2388 | -103.4339
17.2426 | -103.4426
17.2463 | -103.4515
17.2501 | -103.4598
17.2538 | -103.4708
17.2575 | -103.4829
17.2613 | -103.4959
17.265 | -103.507
17.2688 | -103.5178
17.2725 | -103.5322
17.2763 | -103.5511
17.28 | -103.5654
17.2837 | -103.5783
17.2875 | -103.5911
17.2912 | -103.6036
17.295 | -103.617
17.2987 | -103.633
17.3025 | -103.6505
17.3062 | -103.6647
17.3099 | -103.6804
17.3137 | -103.6982
17.3174 | -103.7158
17.3212 | -103.7375
17.3249 | -103.7584
17.3287 | -103.7767
17.3324 | -103.7974
17.3361 | -103.818
17.3399 | -103.8381
17.3436 | -103.8576
17.3474 | -103.8772
17.3511 | -103.8955
17.3549 | -103.913
17.3586 | -103.9304
17.3623 | -103.9484
17.3661 | -103.9648
17.3698 | -103.9793
17.3736 | -103.9931
17.3773 | -104.0063
17.3811 | -104.0211
17.3848 | -104.0376
17.3885 | -104.0488
17.3923 | -104.0571
17.396 | -104.0672
17.3998 | -104.0779
17.4035 | -104.089
17.4073 | -104.0992
17.411 | -104.1074
17.4147 | -104.1133
17.4185 | -104.1205
17.4222 | -104.128
17.426 | -104.1354
17.4297 | -104.1438
17.4335 | -104.1532
17.4372 | -104.1596
17.4409 | -104.1644
17.4447 | -104.1712
17.4484 | -104.1772
17.4522 | -104.1838
17.4559 | -104.1917
17.4596 | -104.2014
17.4634 | -104.2065
17.4671 | -104.2132
17.4709 | -104.2217
17.4746 | -104.233
17.4784 | -104.2396
17.4821 | -104.2422
17.4858 | -104.2461
17.4896 | -104.2512
17.4933 | -104.2589
17.4971 | -104.2672
17.5008 | -104.2756
17.5046 | -104.2835
17.5083 | -104.2883
17.512 | -104.2925
17.5158 | -104.2958
17.5195 | -104.2971
17.5233 | -104.2956
17.527 | -104.2949
17.5308 | -104.2967
17.5345 | -104.2977
17.5382 | -104.2982
17.542 | -104.2983
17.5457 | -104.2999
17.5495 | -104.3
17.5532 | -104.2982
17.557 | -104.2968
17.5607 | -104.2943
17.5644 | -104.2896
17.5682 | -104.2871
17.5719 | -104.2824
17.5757 | -104.2786
17.5794 | -104.278
17.5832 | -104.2753
17.5869 | -104.265
17.5906 | -104.2534
17.5944 | -104.2439
17.5981 | -104.2333
17.6019 | -104.2238
17.6056 | -104.2137
17.6094 | -104.2029
17.6131 | -104.1942
17.6168 | -104.1833
17.6206 | -104.1705
17.6243 | -104.1574
17.6281 | -104.1452
17.6318 | -104.1343
17.6356 | -104.1218
17.6393 | -104.108
17.643 | -104.0948
17.6468 | -104.0803
17.6505 | -104.0652
17.6543 | -104.0515
17.658 | -104.0389
17.6618 | -104.0254
17.6655 | -104.0114
17.6692 | -103.9954
17.673 | -103.9825
17.6767 | -103.9699
17.6805 | -103.9541
17.6842 | -103.9398
17.688 | -103.9266
17.6917 | -103.9156
17.6954 | -103.9042
17.6992 | -103.8913
17.7029 | -103.8756
17.7067 | -103.8618
17.7104 | -103.8464
17.7142 | -103.8317
17.7179 | -103.817
17.7216 | -103.802
17.7254 | -103.7819
17.7291 | -103.7636
17.7329 | -103.7488
17.7366 | -103.7365
17.7404 | -103.7231
17.7441 | -103.7073
17.7478 | -103.6898
17.7516 | -103.6692
17.7553 | -103.651
17.7591 | -103.635
17.7628 | -103.6188
17.7665 | -103.5997
17.7703 | -103.5774
17.774 | -103.558
17.7778 | -103.5402
17.7815 | -103.5179
17.7853 | -103.4957
17.789 | -103.4753
17.7927 | -103.4554
17.7965 | -103.433
17.8002 | -103.4114
17.804 | -103.3896
17.8077 | -103.3661
17.8115 | -103.3446
17.8152 | -103.3294
17.8189 | -103.314
17.8227 | -103.2973
17.8264 | -103.2798
17.8302 | -103.2629
17.8339 | -103.2461
17.8377 | -103.2325
17.8414 | -103.2253
17.8451 | -103.2212
17.8489 | -103.2158
17.8526 | -103.2137
17.8564 | -103.2121
17.8601 | -103.2128
17.8639 | -103.2141
17.8676 | -103.2148
17.8713 | -103.2187
17.8751 | -103.2267
17.8788 | -103.2349
17.8826 | -103.2458
17.8863 | -103.2568
17.8901 | -103.2703
17.8938 | -103.2868
17.8975 | -103.3053
17.9013 | -103.3237
17.905 | -103.3442
17.9088 | -103.3645
17.9125 | -103.3842
17.9163 | -103.4051
17.92 | -103.4289
17.9237 | -103.4527
17.9275 | -103.4744
17.9312 | -103.4968
17.935 | -103.5206
17.9387 | -103.5436
17.9425 | -103.566
17.9462 | -103.5877
17.9499 | -103.6072
17.9537 | -103.6266
17.9574 | -103.6483
17.9612 | -103.6682
17.9649 | -103.6872
17.9687 | -103.7044
17.9724 | -103.7158
17.9761 | -103.7218
17.9799 | -103.7323
17.9836 | -103.7417
17.9874 | -103.747
17.9911 | -103.7457
17.9949 | -103.7391
17.9986 | -103.734
18.0023 | -103.729
18.0061 | -103.7224
18.0098 | -103.7151
18.0136 | -103.7059
18.0173 | -103.6937
18.0211 | -103.6819
18.0248 | -103.669
18.0285 | -103.6501
18.0323 | -103.6277
18.036 | -103.6066
18.0398 | -103.5843
18.0435 | -103.5576
18.0473 | -103.5287
18.051 | -103.5024
18.0547 | -103.4751
18.0585 | -103.4495
18.0622 | -103.4252
18.066 | -103.3988
18.0697 | -103.373
18.0735 | -103.3479
18.0772 | -103.3227
18.0809 | -103.2996
18.0847 | -103.2796
18.0884 | -103.2617
18.0922 | -103.2419
18.0959 | -103.2207
18.0996 | -103.2
18.1034 | -103.1826
18.1071 | -103.168
18.1109 | -103.1518
18.1146 | -103.1383
18.1184 | -103.126
18.1221 | -103.1169
18.1258 | -103.1085
18.1296 | -103.1002
18.1333 | -103.0914
18.1371 | -103.0833
18.1408 | -103.0761
18.1446 | -103.0715
18.1483 | -103.0653
18.152 | -103.0563
18.1558 | -103.0462
18.1595 | -103.0391
18.1633 | -103.0318
18.167 | -103.0219
18.1708 | -103.0082
18.1745 | -102.9961
18.1782 | -102.9837
18.182 | -102.9699
18.1857 | -102.9553
18.1895 | -102.9361
18.1932 | -102.9169
18.197 | -102.899
18.2007 | -102.8797
18.2044 | -102.8577
18.2082 | -102.8344
18.2119 | -102.812
18.2157 | -102.7907
18.2194 | -102.7685
18.2232 | -102.7471
18.2269 | -102.7244
18.2306 | -102.7051
18.2344 | -102.6893
18.2381 | -102.6761
18.2419 | -102.6623
18.2456 | -102.6456
18.2494 | -102.6329
18.2531 | -102.6267
18.2568 | -102.6215
18.2606 | -102.6191
18.2643 | -102.6192
18.2681 | -102.6213
18.2718 | -102.6255
18.2756 | -102.6306
18.2793 | -102.6383
18.283 | -102.6495
18.2868 | -102.6615
18.2905 | -102.6756
18.2943 | -102.6921
18.298 | -102.7087
18.3018 | -102.726
18.3055 | -102.7419
18.3092 | -102.7598
18.313 | -102.7791
18.3167 | -102.796
18.3205 | -102.8118
18.3242 | -102.827
18.328 | -102.8432
18.3317 | -102.8597
18.3354 | -102.8757
18.3392 | -102.8918
18.3429 | -102.9066
18.3467 | -102.9188
18.3504 | -102.9297
18.3542 | -102.9409
18.3579 | -102.9421
18.3616 | -102.951
18.3654 | -102.9569
18.3691 | -102.9601
18.3729 | -102.9655
18.3766 | -102.9701
18.3804 | -102.9738
18.3841 | -102.9762
18.3878 | -102.9792
18.3916 | -102.9813
18.3953 | -102.9826
18.3991 | -102.984
18.4028 | -102.9831
18.4065 | -102.9819
18.4103 | -102.9831
18.414 | -102.9836
18.4178 | -102.9826
18.4215 | -102.9805
18.4253 | -102.9801
18.429 | -102.9822
18.4327 | -102.9876
18.4365 | -102.9889
18.4402 | -102.9881
18.444 | -102.9897
18.4477 | -102.9936
18.4515 | -102.9994
18.4552 | -103.0057
18.4589 | -103.0111
18.4627 | -103.0163
18.4664 | -103.0225
18.4702 | -103.0316
18.4739 | -103.0402
18.4777 | -103.0434
18.4814 | -103.0449
18.4851 | -103.0505
18.4889 | -103.0587
18.4926 | -103.0631
18.4964 | -103.063
18.5001 | -103.0612
18.5039 | -103.0582
18.5076 | -103.0583
18.5113 | -103.0553
18.5151 | -103.0462
18.5188 | -103.036
18.5226 | -103.0244
18.5263 | -103.009
18.5301 | -102.989
18.5338 | -102.9676
18.5375 | -102.948
18.5413 | -102.9291
18.545 | -102.9068
18.5488 | -102.8835
18.5525 | -102.856
18.5563 | -102.8279
18.56 | -102.8006
18.5637 | -102.7751
18.5675 | -102.7523
18.5712 | -102.7254
18.575 | -102.6977
18.5787 | -102.6735
18.5825 | -102.6508
18.5862 | -102.6285
18.5899 | -102.6036
18.5937 | -102.58
18.5974 | -102.5566
18.6012 | -102.5338
18.6049 | -102.5128
18.6087 | -102.496
18.6124 | -102.4842
18.6161 | -102.4723
18.6199 | -102.4573
18.6236 | -102.4418
18.6274 | -102.4301
18.6311 | -102.4202
18.6349 | -102.4149
18.6386 | -102.4082
18.6423 | -102.4024
18.6461 | -102.3997
18.6498 | -102.3994
18.6536 | -102.4
18.6573 | -102.4036
18.6611 | -102.4073
18.6648 | -102.4108
18.6685 | -102.418
18.6723 | -102.4268
18.676 | -102.4361
18.6798 | -102.4482
18.6835 | -102.4624
18.6873 | -102.4759
18.691 | -102.4879
18.6947 | -102.4986
18.6985 | -102.5088
18.7022 | -102.5221
18.706 | -102.5389
18.7097 | -102.5519
18.7135 | -102.5616
18.7172 | -102.5706
18.7209 | -102.5793
18.7247 | -102.5885
18.7284 | -102.6003
18.7322 | -102.6093
18.7359 | -102.6116
18.7396 | -102.6123
18.7434 | -102.6194
18.7471 | -102.6262
18.7509 | -102.6273
18.7546 | -102.6254
18.7584 | -102.6267
18.7621 | -102.6293
18.7658 | -102.6277
18.7696 | -102.6221
18.7733 | -102.6133
18.7771 | -102.6051
18.7808 | -102.5979
18.7846 | -102.5894
18.7883 | -102.5783
18.792 | -102.5641
18.7958 | -102.5474
18.7995 | -102.5304
18.8033 | -102.5141
18.807 | -102.4942
18.8108 | -102.4761
18.8145 | -102.4571
18.8182 | -102.4376
18.822 | -102.415
18.8257 | -102.3937
18.8295 | -102.3731
18.8332 | -102.3529
18.837 | -102.3346
18.8407 | -102.3151
18.8444 | -102.2936
18.8482 | -102.2754
18.8519 | -102.2589
18.8557 | -102.2435
18.8594 | -102.2296
18.8632 | -102.2182
18.8669 | -102.211
18.8706 | -102.2057
18.8744 | -102.201
18.8781 | -102.1966
18.8819 | -102.195
18.8856 | -102.1975
18.8894 | -102.2045
18.8931 | -102.2139
18.8968 | -102.2283
18.9006 | -102.2333
18.9043 | -102.2396
18.9081 | -102.2496
18.9118 | -102.2619
18.9156 | -102.2748
18.9193 | -102.2861
18.923 | -102.2981
18.9268 | -102.3084
18.9305 | -102.3176
18.9343 | -102.3299
18.938 | -102.3404
18.9418 | -102.3499
18.9455 | -102.3572
18.9492 | -102.3624
18.953 | -102.365
18.9567 | -102.3647
18.9605 | -102.3653
18.9642 | -102.3659
18.968 | -102.3653
18.9717 | -102.3587
18.9754 | -102.346
18.9792 | -102.3317
18.9829 | -102.3173
18.9867 | -102.3024
18.9904 | -102.2851
18.9942 | -102.2625
18.9979 | -102.2355
19.0016 | -102.209
19.0054 | -102.1844
19.0091 | -102.1554
19.0129 | -102.1203
19.0166 | -102.0852
19.0204 | -102.0497
19.0241 | -102.012
19.0278 | -101.9705
19.0316 | -101.9286
19.0353 | -101.8893
19.0391 | -101.855
19.0428 | -101.8193
19.0465 | -101.7826
19.0503 | -101.7496
19.054 | -101.7185
19.0578 | -101.689
19.0615 | -101.6603
19.0653 | -101.6338
19.069 | -101.6125
19.0727 | -101.5919
19.0765 | -101.574
19.0802 | -101.5547
19.084 | -101.5405
19.0877 | -101.5308
19.0915 | -101.5236
19.0952 | -101.5171
19.0989 | -101.5101
19.1027 | -101.5042
19.1064 | -101.5038
19.1102 | -101.5079
19.1139 | -101.5157
19.1177 | -101.5235
19.1214 | -101.5342
19.1251 | -101.5464
19.1289 | -101.5584
19.1326 | -101.573
19.1364 | -101.5857
19.1401 | -101.6003
19.1439 | -101.6169
19.1476 | -101.6345
19.1513 | -101.6522
19.1551 | -101.6721
19.1588 | -101.6919
19.1626 | -101.7108
19.1663 | -101.7324
19.1701 | -101.7581
19.1738 | -101.7818
19.1775 | -101.8068
19.1813 | -101.8311
19.185 | -101.8513
19.1888 | -101.8722
19.1925 | -101.894
19.1963 | -101.9166
19.2 | -101.944
19.2037 | -101.9735
19.2075 | -101.9992
19.2112 | -102.0248
19.215 | -102.0505
19.2187 | -102.0786
19.2225 | -102.1094
19.2262 | -102.1413
19.2299 | -102.1725
19.2337 | -102.2
19.2374 | -102.2275
19.2412 | -102.2536
19.2449 | -102.2777
19.2487 | -102.3006
19.2524 | -102.3268
19.2561 | -102.3503
19.2599 | -102.3683
19.2636 | -102.3834
19.2674 | -102.3976
19.2711 | -102.4097
19.2749 | -102.4161
19.2786 | -102.4217
19.2823 | -102.4312
19.2861 | -102.4373
19.2898 | -102.4353
19.2936 | -102.4281
19.2973 | -102.4242
19.3011 | -102.4226
19.3048 | -102.4188
19.3085 | -102.4081
19.3123 | -102.3913
19.316 | -102.3734
19.3198 | -102.3554
19.3235 | -102.3334
19.3273 | -102.3095
19.331 | -102.2856
19.3347 | -102.2639
19.3385 | -102.2395
19.3422 | -102.2096
19.346 | -102.181
19.3497 | -102.1538
19.3534 | -102.1292
19.3572 | -102.1037
19.3609 | -102.0789
19.3647 | -102.0558
19.3684 | -102.034
19.3722 | -102.0168
19.3759 | -102.0023
19.3796 | -101.9876
19.3834 | -101.9759
19.3871 | -101.967
19.3909 | -101.9622
19.3946 | -101.9608
19.3984 | -101.9619
19.4021 | -101.9665
19.4058 | -101.9742
19.4096 | -101.9874
19.4133 | -102.003
19.4171 | -102.0206
19.4208 | -102.0399
19.4246 | -102.0651
19.4283 | -102.0945
19.432 | -102.1262
19.4358 | -102.1784
19.4395 | -102.215
19.4433 | -102.2525
19.447 | -102.2942
19.4508 | -102.3386
19.4545 | -102.3793
19.4582 | -102.4211
19.462 | -102.4681
19.4657 | -102.5163
19.4695 | -102.5636
19.4732 | -102.6061
19.477 | -102.6436
19.4807 | -102.6827
19.4844 | -102.7239
19.4882 | -102.761
19.4919 | -102.7976
19.4957 | -102.8323
19.4994 | -102.862
19.5032 | -102.8894
19.5069 | -102.9146
19.5106 | -102.9374
19.5144 | -102.9554
19.5181 | -102.9716
19.5219 | -102.9904
19.5256 | -103.0056
19.5294 | -103.0135
19.5331 | -103.0182
19.5368 | -103.0235
19.5406 | -103.0275
19.5443 | -103.03
19.5481 | -103.0287
19.5518 | -103.0261
19.5556 | -103.0216
19.5593 | -103.0116
19.563 | -103.0019
19.5668 | -102.9947
19.5705 | -102.9888
19.5743 | -102.9749
19.578 | -102.9554
19.5818 | -102.9374
19.5855 | -102.9219
19.5892 | -102.9097
19.593 | -102.9006
19.5967 | -102.8898
19.6005 | -102.8773
19.6042 | -102.863
19.608 | -102.8502
19.6117 | -102.839
19.6154 | -102.8304
19.6192 | -102.8217
19.6229 | -102.8106
19.6267 | -102.7987
19.6304 | -102.7897
19.6342 | -102.7789
19.6379 | -102.7664
19.6416 | -102.7558
19.6454 | -102.7455
19.6491 | -102.7326
19.6529 | -102.7204
19.6566 | -102.7067
19.6604 | -102.692
19.6641 | -102.6783
19.6678 | -102.6672
19.6716 | -102.6558
19.6753 | -102.647
19.6791 | -102.6367
19.6828 | -102.6245
19.6866 | -102.6135
19.6903 | -102.6005
19.694 | -102.588
19.6978 | -102.581
19.7015 | -102.571
19.7053 | -102.5588
19.709 | -102.5475
19.7127 | -102.5371
19.7165 | -102.5254
19.7202 | -102.5144
19.724 | -102.5054
19.7277 | -102.4949
19.7315 | -102.4839
19.7352 | -102.4765
19.7389 | -102.4718
19.7427 | -102.4674
19.7464 | -102.4622
19.7502 | -102.4579
19.7539 | -102.4546
19.7577 | -102.4499
19.7614 | -102.4449
19.7651 | -102.4423
19.7689 | -102.4422
19.7726 | -102.4417
19.7764 | -102.4394
19.7801 | -102.4395
19.7839 | -102.4392
19.7876 | -102.4413
19.7913 | -102.446
19.7951 | -102.4508
19.7988 | -102.4546
19.8026 | -102.4591
19.8063 | -102.4664
19.8101 | -102.4712
19.8138 | -102.4742
19.8175 | -102.4778
19.8213 | -102.4822
19.825 | -102.4873
19.8288 | -102.4925
19.8325 | -102.498
19.8363 | -102.499
19.84 | -102.5006
19.8437 | -102.5012
19.8475 | -102.4995
19.8512 | -102.4939
19.855 | -102.4896
19.8587 | -102.4856
19.8625 | -102.4773
19.8662 | -102.4681
19.8699 | -102.4553
19.8737 | -102.4374
19.8774 | -102.4161
19.8812 | -102.3966
19.8849 | -102.378
19.8887 | -102.3583
19.8924 | -102.3369
19.8961 | -102.311
19.8999 | -102.2852
19.9036 | -102.2609
19.9074 | -102.235
19.9111 | -102.2054
19.9149 | -102.1752
19.9186 | -102.1463
19.9223 | -102.1161
19.9261 | -102.0851
19.9298 | -102.054
19.9336 | -102.0186
19.9373 | -101.9846
19.9411 | -101.9531
19.9448 | -101.9209
19.9485 | -101.8874
19.9523 | -101.8551
19.956 | -101.8224
19.9598 | -101.7891
19.9635 | -101.7594
19.9673 | -101.7322
19.971 | -101.7078
19.9747 | -101.6952
19.9785 | -101.6734
19.9822 | -101.6515
19.986 | -101.6319
19.9897 | -101.6146
19.9934 | -101.5993
19.9972 | -101.5856
20.0009 | -101.5739
20.0047 | -101.5628
20.0084 | -101.5545
20.0122 | -101.5483
20.0159 | -101.5427
20.0196 | -101.5388
20.0234 | -101.5351
20.0271 | -101.5304
20.0309 | -101.5285
20.0346 | -101.5315
20.0384 | -101.5382
20.0421 | -101.5452
20.0458 | -101.5501
20.0496 | -101.5547
20.0533 | -101.5621
20.0571 | -101.5713
20.0608 | -101.5793
20.0646 | -101.5861
20.0683 | -101.5955
20.072 | -101.6063
20.0758 | -101.6156
20.0795 | -101.6267
20.0833 | -101.6402
20.087 | -101.6501
20.0908 | -101.6613
20.0945 | -101.6753
20.0982 | -101.6915
20.102 | -101.7067
20.1057 | -101.7214
20.1095 | -101.7354
20.1132 | -101.7496
20.117 | -101.7631
20.1207 | -101.7785
20.1244 | -101.7909
20.1282 | -101.8038
20.1319 | -101.8189
20.1357 | -101.8345
20.1394 | -101.8475
20.1432 | -101.8596
20.1469 | -101.8717
20.1506 | -101.8817
20.1544 | -101.89
20.1581 | -101.9028
20.1619 | -101.915
20.1656 | -101.925
20.1694 | -101.9354
20.1731 | -101.9435
20.1768 | -101.9507
20.1806 | -101.9599
20.1843 | -101.9686
20.1881 | -101.9758
20.1918 | -101.9835
20.1956 | -101.9941
20.1993 | -102.0067
20.203 | -102.0183
20.2068 | -102.0318
20.2105 | -102.0445
20.2143 | -102.0589
20.218 | -102.0764
20.2218 | -102.0915
20.2255 | -102.1041
20.2292 | -102.1179
20.233 | -102.131
20.2367 | -102.1461
20.2405 | -102.1635
20.2442 | -102.1822
20.248 | -102.197
20.2517 | -102.2094
20.2554 | -102.2241
20.2592 | -102.236
20.2629 | -102.2499
20.2667 | -102.2634
20.2704 | -102.2761
20.2742 | -102.2861
20.2779 | -102.2973
20.2816 | -102.3075
20.2854 | -102.3145
20.2891 | -102.3241
20.2929 | -102.3339
20.2966 | -102.3428
20.3004 | -102.3508
20.3041 | -102.3592
20.3078 | -102.3648
20.3116 | -102.3705
20.3153 | -102.3795
20.3191 | -102.3855
20.3228 | -102.388
20.3266 | -102.3882
20.3303 | -102.3897
20.334 | -102.3934
20.3378 | -102.3943
20.3415 | -102.3928
20.3453 | -102.3865
20.349 | -102.3799
20.3527 | -102.3757
20.3565 | -102.3696
20.3602 | -102.3583
20.364 | -102.3437
20.3677 | -102.3246
20.3715 | -102.3086
20.3752 | -102.2941
20.3789 | -102.2784
20.3827 | -102.2598
20.3864 | -102.2385
20.3902 | -102.2167
20.3939 | -102.1925
20.3977 | -102.1698
20.4014 | -102.1454
20.4051 | -102.122
20.4089 | -102.0955
20.4126 | -102.0686
20.4164 | -102.0391
20.4201 | -102.0092
20.4239 | -101.9782
20.4276 | -101.9478
20.4313 | -101.9184
20.4351 | -101.8864
20.4388 | -101.8545
20.4426 | -101.8234
20.4463 | -101.789
20.4501 | -101.7549
20.4538 | -101.7219
20.4575 | -101.69
20.4613 | -101.6609
20.465 | -101.6347
20.4688 | -101.6053
20.4725 | -101.5734
20.4763 | -101.546
20.48 | -101.5224
20.4837 | -101.4996
20.4875 | -101.4772
20.4912 | -101.4543
20.495 | -101.4307
20.4987 | -101.4099
20.5025 | -101.3952
20.5062 | -101.3826
20.5099 | -101.3732
20.5137 | -101.3653
20.5174 | -101.3692
20.5212 | -101.3635
20.5249 | -101.361
20.5287 | -101.3582
20.5324 | -101.3564
20.5361 | -101.3573
20.5399 | -101.3604
20.5436 | -101.3635
20.5474 | -101.3654
20.5511 | -101.3702
20.5549 | -101.3748
20.5586 | -101.3826
20.5623 | -101.3906
20.5661 | -101.4004
20.5698 | -101.4098
20.5736 | -101.4189
20.5773 | -101.4284
20.5811 | -101.4359
20.5848 | -101.4449
20.5885 | -101.4534
20.5923 | -101.4623
20.596 | -101.4708
20.5998 | -101.4792
20.6035 | -101.4849
20.6073 | -101.4869
20.611 | -101.4875
20.6147 | -101.4908
20.6185 | -101.4963
20.6222 | -101.5002
20.626 | -101.5021
20.6297 | -101.506
20.6334 | -101.5097
20.6372 | -101.5142
20.6409 | -101.5195
20.6447 | -101.522
20.6484 | -101.5223
20.6522 | -101.5248
20.6559 | -101.5261
20.6596 | -101.5266
20.6634 | -101.5235
20.6671 | -101.5221
20.6709 | -101.5186
20.6746 | -101.5149
20.6784 | -101.507
20.6821 | -101.5005
20.6858 | -101.4974
20.6896 | -101.4932
20.6933 | -101.491
20.6971 | -101.488
20.7008 | -101.4863
20.7046 | -101.4844
20.7083 | -101.4826
20.712 | -101.4798
20.7158 | -101.477
20.7195 | -101.4749
20.7233 | -101.4746
20.727 | -101.4772
20.7308 | -101.4782
20.7345 | -101.4796
20.7382 | -101.4796
20.742 | -101.4842
20.7457 | -101.4887
20.7495 | -101.4923
20.7532 | -101.4956
20.757 | -101.4982
20.7607 | -101.5032
20.7644 | -101.5075
20.7682 | -101.5106
20.7719 | -101.515
20.7757 | -101.5197
20.7794 | -101.5262
20.7832 | -101.5336
20.7869 | -101.5402
20.7906 | -101.5456
20.7944 | -101.5428
20.7981 | -101.5418
20.8019 | -101.5399
20.8056 | -101.5384
20.8094 | -101.5358
20.8131 | -101.5316
20.8168 | -101.5266
20.8206 | -101.519
20.8243 | -101.5114
20.8281 | -101.5051
20.8318 | -101.4965
20.8356 | -101.4841
20.8393 | -101.4693
20.843 | -101.4553
20.8468 | -101.4397
20.8505 | -101.4241
20.8543 | -101.4052
20.858 | -101.3843
20.8618 | -101.3638
20.8655 | -101.3417
20.8692 | -101.3183
20.873 | -101.2958
20.8767 | -101.2735
20.8805 | -101.2512
20.8842 | -101.2286
20.888 | -101.2064
20.8917 | -101.1843
20.8954 | -101.1635
20.8992 | -101.1453
20.9029 | -101.1289
20.9067 | -101.1132
20.9104 | -101.0964
20.9142 | -101.0796
20.9179 | -101.0651
20.9216 | -101.0537
20.9254 | -101.0416
20.9291 | -101.0333
20.9329 | -101.0294
20.9366 | -101.0261
20.9404 | -101.0245
20.9441 | -101.025
20.9478 | -101.0258
20.9516 | -101.0329
20.9553 | -101.0417
20.9591 | -101.0516
20.9628 | -101.0621
20.9666 | -101.0745
20.9703 | -101.0909
20.974 | -101.1124
20.9778 | -101.1314
20.9815 | -101.1469
20.9853 | -101.1631
20.989 | -101.1823
20.9927 | -101.2035
20.9965 | -101.2261
21.0002 | -101.249
21.004 | -101.2718
21.0077 | -101.296
21.0115 | -101.3215
21.0152 | -101.3459
21.0189 | -101.3679
21.0227 | -101.3884
21.0264 | -101.409
21.0302 | -101.4286
21.0339 | -101.4488
21.0377 | -101.4638
21.0414 | -101.4746
21.0451 | -101.4869
21.0489 | -101.501
21.0526 | -101.5144
21.0564 | -101.5181
21.0601 | -101.5305
21.0639 | -101.5436
21.0676 | -101.5544
21.0713 | -101.5614
21.0751 | -101.5685
21.0788 | -101.5742
21.0826 | -101.5804
21.0863 | -101.5856
21.0901 | -101.5888
21.0938 | -101.5944
21.0975 | -101.5967
21.1013 | -101.5989
21.105 | -101.6001
21.1088 | -101.6021
21.1125 | -101.6038
21.1163 | -101.6048
21.12 | -101.6041
21.1237 | -101.6041
21.1275 | -101.6024
21.1312 | -101.6018
21.135 | -101.6054
21.1387 | -101.6104
21.1425 | -101.6124
21.1462 | -101.614
21.1499 | -101.6147
21.1537 | -101.6159
21.1574 | -101.6172
21.1612 | -101.6187
21.1649 | -101.6202
21.1687 | -101.6244
21.1724 | -101.6281
21.1761 | -101.6298
21.1799 | -101.6294
21.1836 | -101.6269
21.1874 | -101.6244
21.1911 | -101.6267
21.1949 | -101.6291
21.1986 | -101.6308
21.2023 | -101.6355
21.2061 | -101.6415
21.2098 | -101.648
21.2136 | -101.6533
21.2173 | -101.6564
21.2211 | -101.6593
21.2248 | -101.6644
21.2285 | -101.666
21.2323 | -101.664
21.236 | -101.6655
21.2398 | -101.6698
21.2435 | -101.6735
21.2473 | -101.6765
21.251 | -101.6762
21.2547 | -101.6753
21.2585 | -101.6738
21.2622 | -101.6722
21.266 | -101.6713
21.2697 | -101.668
21.2735 | -101.6634
21.2772 | -101.6576
21.2809 | -101.6547
21.2847 | -101.6511
21.2884 | -101.6461
21.2922 | -101.6386
21.2959 | -101.6287
21.2996 | -101.6206
21.3034 | -101.6127
21.3071 | -101.6014
21.3109 | -101.5898
21.3146 | -101.5778
21.3184 | -101.5662
21.3221 | -101.5551
21.3258 | -101.5448
21.3296 | -101.5356
21.3333 | -101.5258
21.3371 | -101.5136
21.3408 | -101.5033
21.3446 | -101.4909
21.3483 | -101.4758
21.352 | -101.4624
21.3558 | -101.4478
21.3595 | -101.4331
21.3633 | -101.4203
21.367 | -101.4081
21.3708 | -101.3957
21.3745 | -101.3845
21.3782 | -101.3732
21.382 | -101.3624
21.3857 | -101.3536
21.3895 | -101.3486
21.3932 | -101.3433
21.397 | -101.3367
21.4007 | -101.332
21.4044 | -101.326
21.4082 | -101.3202
21.4119 | -101.3145
21.4157 | -101.3083
21.4194 | -101.3038
21.4232 | -101.3014
21.4269 | -101.299
21.4306 | -101.298
21.4344 | -101.2997
21.4381 | -101.3021
21.4419 | -101.304
21.4456 | -101.3059
21.4494 | -101.3058
21.4531 | -101.305
21.4568 | -101.3047
21.4606 | -101.3036
21.4643 | -101.3038
21.4681 | -101.3031
21.4718 | -101.3012
21.4756 | -101.2966
21.4793 | -101.2956
21.483 | -101.2967
21.4868 | -101.2981
21.4905 | -101.2972
21.4943 | -101.2918
21.498 | -101.2849
21.5018 | -101.2776
21.5055 | -101.265
21.5092 | -101.2501
21.513 | -101.2356
21.5167 | -101.2212
21.5205 | -101.2078
21.5242 | -101.1902
21.528 | -101.1693
21.5317 | -101.1481
21.5354 | -101.1276
21.5392 | -101.1052
21.5429 | -101.0797
21.5467 | -101.0537
21.5504 | -101.0281
21.5542 | -101.0028
21.5579 | -100.9802
21.5616 | -100.9556
21.5654 | -100.9325
21.5691 | -100.91
21.5729 | -100.887
21.5766 | -100.8666
21.5804 | -100.849
21.5841 | -100.8309
21.5878 | -100.8137
21.5916 | -100.7972
21.5953 | -100.7982
21.5991 | -100.7874
21.6028 | -100.7799
21.6065 | -100.774
21.6103 | -100.7698
21.614 | -100.7679
21.6178 | -100.7658
21.6215 | -100.7657
21.6253 | -100.7687
21.629 | -100.773
21.6327 | -100.7787
21.6365 | -100.7864
21.6402 | -100.7959
21.644 | -100.8095
21.6477 | -100.8253
21.6515 | -100.8417
21.6552 | -100.8598
21.6589 | -100.8791
21.6627 | -100.9006
21.6664 | -100.9227
21.6702 | -100.9437
21.6739 | -100.9657
21.6777 | -100.9881
21.6814 | -101.0098
21.6851 | -101.0321
21.6889 | -101.0554
21.6926 | -101.0778
21.6964 | -101.0974
21.7001 | -101.1188
21.7039 | -101.1394
21.7076 | -101.1592
21.7113 | -101.1808
21.7151 | -101.1997
21.7188 | -101.2181
21.7226 | -101.2353
21.7263 | -101.249
21.7301 | -101.2612
21.7338 | -101.271
21.7375 | -101.275
21.7413 | -101.2724
21.745 | -101.269
21.7488 | -101.2667
21.7525 | -101.263
21.7563 | -101.2598
21.76 | -101.2562
21.7637 | -101.2488
21.7675 | -101.2389
21.7712 | -101.2237
21.775 | -101.2062
21.7787 | -101.1887
21.7825 | -101.1718
21.7862 | -101.1494
21.7899 | -101.1244
21.7937 | -101.0962
21.7974 | -101.0647
21.8012 | -101.0337
21.8049 | -101.0023
21.8087 | -100.9716
21.8124 | -100.9421
21.8161 | -100.9113
21.8199 | -100.8773
21.8236 | -100.843
21.8274 | -100.8086
21.8311 | -100.777
21.8349 | -100.7452
21.8386 | -100.7121
21.8423 | -100.68
21.8461 | -100.6477
21.8498 | -100.6136
21.8536 | -100.5799
21.8573 | -100.5502
21.8611 | -100.522
21.8648 | -100.4938
21.8685 | -100.4692
21.8723 | -100.4416
21.876 | -100.4134
21.8798 | -100.3893
21.8835 | -100.3654
21.8873 | -100.3426
21.891 | -100.3223
21.8947 | -100.3043
21.8985 | -100.2873
21.9022 | -100.2744
21.906 | -100.263
21.9097 | -100.2538
21.9135 | -100.245
21.9172 | -100.237
21.9209 | -100.2315
21.9247 | -100.2272
21.9284 | -100.2254
21.9322 | -100.2243
21.9359 | -100.2231
21.9396 | -100.2233
21.9434 | -100.223
21.9471 | -100.2233
21.9509 | -100.2264
21.9546 | -100.2281
21.9584 | -100.2298
21.9621 | -100.2319
21.9658 | -100.2349
21.9696 | -100.2389
21.9733 | -100.2444
21.9771 | -100.2508
21.9808 | -100.2578
21.9846 | -100.2642
21.9883 | -100.2702
21.992 | -100.2768
21.9958 | -100.2818
21.9995 | -100.2847
22.0033 | -100.2881
22.007 | -100.2941
22.0108 | -100.2945
22.0145 | -100.2941
22.0182 | -100.2956
22.022 | -100.297
22.0257 | -100.3003
22.0295 | -100.3025
22.0332 | -100.3019
22.037 | -100.3013
22.0407 | -100.3016
22.0444 | -100.3018
22.0482 | -100.3033
22.0519 | -100.3059
22.0557 | -100.3054
22.0594 | -100.3029
22.0632 | -100.3028
22.0669 | -100.3021
22.0706 | -100.3
22.0744 | -100.2966
22.0781 | -100.2924
22.0819 | -100.2882
22.0856 | -100.2834
22.0894 | -100.2762
22.0931 | -100.2674
22.0968 | -100.258
22.1006 | -100.25
22.1043 | -100.2419
22.1081 | -100.2358
22.1118 | -100.2308
22.1156 | -100.2241
22.1193 | -100.216
22.123 | -100.2095
22.1268 | -100.2031
22.1305 | -100.1964
22.1343 | -100.1896
22.138 | -100.1853
22.1418 | -100.1799
22.1455 | -100.1739
22.1492 | -100.1698
22.153 | -100.1678
22.1567 | -100.1651
22.1605 | -100.1633
22.1642 | -100.1623
22.168 | -100.163
22.1717 | -100.1607
22.1754 | -100.1584
22.1792 | -100.1574
22.1829 | -100.1574
22.1867 | -100.1629
22.1904 | -100.1702
22.1942 | -100.1758
22.1979 | -100.1832
22.2016 | -100.1929
22.2054 | -100.2023
22.2091 | -100.212
22.2129 | -100.2227
22.2166 | -100.2321
22.2204 | -100.2417
22.2241 | -100.2531
22.2278 | -100.2666
22.2316 | -100.2809
22.2353 | -100.2968
22.2391 | -100.3136
22.2428 | -100.3294
22.2465 | -100.3443
22.2503 | -100.3604
22.254 | -100.3792
22.2578 | -100.4001
22.2615 | -100.4215
22.2653 | -100.4404
22.269 | -100.4598
22.2727 | -100.4779
22.2765 | -100.4957
22.2802 | -100.5111
22.284 | -100.5227
22.2877 | -100.5323
22.2915 | -100.5419
22.2952 | -100.5528
22.2989 | -100.5638
22.3027 | -100.5742
22.3064 | -100.5851
22.3102 | -100.5954
22.3139 | -100.6068
22.3177 | -100.6152
22.3214 | -100.6223
22.3251 | -100.6322
22.3289 | -100.6394
22.3326 | -100.6412
22.3364 | -100.6414
22.3401 | -100.6415
22.3439 | -100.6403
22.3476 | -100.6418
22.3513 | -100.6417
22.3551 | -100.6385
22.3588 | -100.6329
22.3626 | -100.6291
22.3663 | -100.6266
22.3701 | -100.6263
22.3738 | -100.6262
22.3775 | -100.6257
22.3813 | -100.6251
22.385 | -100.6245
22.3888 | -100.6256
22.3925 | -100.6281
22.3963 | -100.6324
22.4 | -100.6362
22.4037 | -100.6408
22.4075 | -100.6461
22.4112 | -100.6515
22.415 | -100.6556
22.4187 | -100.6626
22.4225 | -100.6686
22.4262 | -100.6696
22.4299 | -100.6721
22.4337 | -100.6759
22.4374 | -100.6777
22.4412 | -100.6813
22.4449 | -100.6872
22.4487 | -100.6936
22.4524 | -100.7005
22.4561 | -100.7071
22.4599 | -100.7143
22.4636 | -100.7216
22.4674 | -100.725
22.4711 | -100.7271
22.4749 | -100.7304
22.4786 | -100.7332
22.4823 | -100.7396
22.4861 | -100.7444
22.4898 | -100.751
22.4936 | -100.7591
22.4973 | -100.7657
22.5011 | -100.7694
22.5048 | -100.7708
22.5085 | -100.7717
22.5123 | -100.774
22.516 | -100.7765
22.5198 | -100.7766
22.5235 | -100.7744
22.5273 | -100.7726
22.531 | -100.77
22.5347 | -100.7662
22.5385 | -100.763
22.5422 | -100.758
22.546 | -100.7515
22.5497 | -100.7433
22.5535 | -100.7327
22.5572 | -100.7203
22.5609 | -100.7082
22.5647 | -100.6944
22.5684 | -100.6781
22.5722 | -100.659
22.5759 | -100.6403
22.5796 | -100.6201
22.5834 | -100.597
22.5871 | -100.5728
22.5909 | -100.5498
22.5946 | -100.5265
22.5984 | -100.5027
22.6021 | -100.4782
22.6058 | -100.4557
22.6096 | -100.4333
22.6133 | -100.4108
22.6171 | -100.3869
22.6208 | -100.3613
22.6246 | -100.3344
22.6283 | -100.3085
22.632 | -100.281
22.6358 | -100.2556
22.6395 | -100.2298
22.6433 | -100.2076
22.647 | -100.1872
22.6508 | -100.1678
22.6545 | -100.148
22.6582 | -100.1266
22.662 | -100.1079
22.6657 | -100.0887
22.6695 | -100.0693
22.6732 | -100.0524
22.677 | -100.0427
22.6807 | -100.0317
22.6844 | -100.0192
22.6882 | -100.0057
22.6919 | -99.9897
22.6957 | -99.9754
22.6994 | -99.9617
22.7032 | -99.9487
22.7069 | -99.9371
22.7106 | -99.9274
22.7144 | -99.9191
22.7181 | -99.909
22.7219 | -99.9009
22.7256 | -99.8929
22.7294 | -99.8851
22.7331 | -99.8784
22.7368 | -99.8729
22.7406 | -99.8687
22.7443 | -99.8658
22.7481 | -99.8663
22.7518 | -99.8685
22.7556 | -99.8705
22.7593 | -99.8735
22.763 | -99.8768
22.7668 | -99.8783
22.7705 | -99.8804
22.7743 | -99.8837
22.778 | -99.8871
22.7818 | -99.892
22.7855 | -99.8965
22.7892 | -99.9026
22.793 | -99.9077
22.7967 | -99.9115
22.8005 | -99.9145
22.8042 | -99.9163
22.808 | -99.9203
22.8117 | -99.923
22.8154 | -99.9265
22.8192 | -99.9328
22.8229 | -99.9406
22.8267 | -99.9483
22.8304 | -99.9566
22.8342 | -99.9646
22.8379 | -99.9709
22.8416 | -99.9753
22.8454 | -99.9806
22.8491 | -99.986
22.8529 | -99.9902
22.8566 | -99.9936
22.8604 | -99.9969
22.8641 | -100.0002
22.8678 | -100.0018
22.8716 | -100.0045
22.8753 | -100.0082
22.8791 | -100.0115
22.8828 | -100.0136
22.8865 | -100.0165
22.8903 | -100.0182
22.894 | -100.0195
22.8978 | -100.0173
22.9015 | -100.0154
22.9053 | -100.0154
22.909 | -100.0174
22.9127 | -100.0178
22.9165 | -100.0184
22.9202 | -100.0201
22.924 | -100.0228
22.9277 | -100.0227
22.9315 | -100.0219
22.9352 | -100.0219
22.9389 | -100.0233
22.9427 | -100.0257
22.9464 | -100.0287
22.9502 | -100.0327
22.9539 | -100.0387
22.9577 | -100.0427
22.9614 | -100.048
22.9651 | -100.0546
22.9689 | -100.0622
22.9726 | -100.0694
22.9764 | -100.0798
22.9801 | -100.0944
22.9839 | -100.1089
22.9876 | -100.1228
22.9913 | -100.1377
22.9951 | -100.153
22.9988 | -100.1672
23.0026 | -100.1817
23.0063 | -100.1971
23.0101 | -100.2142
23.0138 | -100.2331
23.0175 | -100.2546
23.0213 | -100.2725
23.025 | -100.2872
23.0288 | -100.2996
23.0325 | -100.313
23.0363 | -100.3272
23.04 | -100.341
23.0437 | -100.3557
23.0475 | -100.3709
23.0512 | -100.3848
23.055 | -100.398
23.0587 | -100.4083
23.0625 | -100.4168
23.0662 | -100.4226
23.0699 | -100.4262
23.0737 | -100.4273
23.0774 | -100.4245
23.0812 | -100.4232
23.0849 | -100.4212
23.0887 | -100.4193
23.0924 | -100.4175
23.0961 | -100.4147
23.0999 | -100.408
23.1036 | -100.3999
23.1074 | -100.3911
23.1111 | -100.3827
23.1149 | -100.3734
23.1186 | -100.3612
23.1223 | -100.3482
23.1261 | -100.3344
23.1298 | -100.3192
23.1336 | -100.3023
23.1373 | -100.2856
23.1411 | -100.2668
23.1448 | -100.2464
23.1485 | -100.2246
23.1523 | -100.2034
23.156 | -100.1829
23.1598 | -100.1609
23.1635 | -100.1398
23.1673 | -100.1223
23.171 | -100.1025
23.1747 | -100.0799
23.1785 | -100.0586
23.1822 | -100.0373
23.186 | -100.0141
23.1897 | -99.9923
23.1934 | -99.9739
23.1972 | -99.9546
23.2009 | -99.9338
23.2047 | -99.9154
23.2084 | -99.899
23.2122 | -99.8831
23.2159 | -99.8721
23.2196 | -99.8568
23.2234 | -99.8423
23.2271 | -99.8317
23.2309 | -99.8207
23.2346 | -99.8101
23.2384 | -99.8002
23.2421 | -99.7919
23.2458 | -99.7848
23.2496 | -99.7787
23.2533 | -99.7741
23.2571 | -99.7702
23.2608 | -99.7668
23.2646 | -99.7638
23.2683 | -99.7626
23.272 | -99.7644
23.2758 | -99.7657
23.2795 | -99.7677
23.2833 | -99.7706
23.287 | -99.7771
23.2908 | -99.7851
23.2945 | -99.792
23.2982 | -99.7984
23.302 | -99.8055
23.3057 | -99.8124
23.3095 | -99.8177
23.3132 | -99.8235
23.317 | -99.8334
23.3207 | -99.8441
23.3244 | -99.8554
23.3282 | -99.8672
23.3319 | -99.8793
23.3357 | -99.8907
23.3394 | -99.9029
23.3432 | -99.9148
23.3469 | -99.9252
23.3506 | -99.9383
23.3544 | -99.9536
23.3581 | -99.9707
23.3619 | -99.9845
23.3656 | -99.9954
23.3694 | -100.0029
23.3731 | -100.0086
23.3768 | -100.0136
23.3806 | -100.0205
23.3843 | -100.0272
23.3881 | -100.0359
23.3918 | -100.0438
23.3956 | -100.052
23.3993 | -100.0611
23.403 | -100.0705
23.4068 | -100.0803
23.4105 | -100.0877
23.4143 | -100.0946
23.418 | -100.1013
23.4218 | -100.1074
23.4255 | -100.1134
23.4292 | -100.1225
23.433 | -100.1309
23.4367 | -100.1375
23.4405 | -100.1444
23.4442 | -100.1524
23.448 | -100.1638
23.4517 | -100.1744
23.4554 | -100.1846
23.4592 | -100.1958
23.4629 | -100.2055
23.4667 | -100.2137
23.4704 | -100.2222
23.4742 | -100.2316
23.4779 | -100.2426
23.4816 | -100.2537
23.4854 | -100.2617
23.4891 | -100.2705
23.4929 | -100.2819
23.4966 | -100.2924
23.5004 | -100.3056
23.5041 | -100.3208
23.5078 | -100.3353
23.5116 | -100.3497
23.5153 | -100.3647
23.5191 | -100.3753
23.5228 | -100.3803
23.5266 | -100.3867
23.5303 | -100.3949
23.534 | -100.4023
23.5378 | -100.4085
23.5415 | -100.4143
23.5453 | -100.4203
23.549 | -100.4261
23.5527 | -100.4297
23.5565 | -100.4325
23.5602 | -100.4323
23.564 | -100.4308
23.5677 | -100.4305
23.5715 | -100.4282
23.5752 | -100.427
23.5789 | -100.4216
23.5827 | -100.4151
23.5864 | -100.4083
23.5902 | -100.4039
23.5939 | -100.3985
23.5977 | -100.3923
23.6014 | -100.3845
23.6051 | -100.3754
23.6089 | -100.3678
23.6126 | -100.3569
23.6164 | -100.3447
23.6201 | -100.3312
23.6239 | -100.3165
23.6276 | -100.3017
23.6313 | -100.2897
23.6351 | -100.2791
23.6388 | -100.2677
23.6426 | -100.2579
23.6463 | -100.2479
23.6501 | -100.2353
23.6538 | -100.2259
23.6575 | -100.2168
23.6613 | -100.2071
23.665 | -100.1982
23.6688 | -100.1907
23.6725 | -100.1829
23.6763 | -100.1747
23.68 | -100.1674
23.6837 | -100.1604
23.6875 | -100.1554
23.6912 | -100.1514
23.695 | -100.15
23.6987 | -100.1487
23.7025 | -100.1485
23.7062 | -100.1487
23.7099 | -100.149
23.7137 | -100.1473
23.7174 | -100.145
23.7212 | -100.1422
23.7249 | -100.1407
23.7287 | -100.1407
23.7324 | -100.1403
23.7361 | -100.1398
23.7399 | -100.1401
23.7436 | -100.1401
23.7474 | -100.1403
23.7511 | -100.1375
23.7549 | -100.1257
23.7586 | -100.1219
23.7623 | -100.1179
23.7661 | -100.1135
23.7698 | -100.1106
23.7736 | -100.1074
23.7773 | -100.1027
23.7811 | -100.0961
23.7848 | -100.0877
23.7885 | -100.0805
23.7923 | -100.0735
23.796 | -100.0648
23.7998 | -100.0568
23.8035 | -100.0493
23.8073 | -100.0405
23.811 | -100.0327
23.8147 | -100.0267
23.8185 | -100.0201
23.8222 | -100.0166
23.826 | -100.0108
23.8297 | -99.9985
23.8334 | -99.9871
23.8372 | -99.9769
23.8409 | -99.9678
23.8447 | -99.9591
23.8484 | -99.9501
23.8522 | -99.9413
23.8559 | -99.9356
23.8596 | -99.9336
23.8634 | -99.9303
23.8671 | -99.9262
23.8709 | -99.924
23.8746 | -99.9231
23.8784 | -99.9228
23.8821 | -99.9233
23.8858 | -99.9236
23.8896 | -99.9265
23.8933 | -99.931
23.8971 | -99.9351
23.9008 | -99.9392
23.9046 | -99.9418
23.9083 | -99.9468
23.912 | -99.95
23.9158 | -99.9554
23.9195 | -99.9617
23.9233 | -99.9678
23.927 | -99.973
23.9308 | -99.9779
23.9345 | -99.9823
23.9382 | -99.9884
23.942 | -99.9958
23.9457 | -100.0032
23.9495 | -100.0132
23.9532 | -100.0228
23.957 | -100.0315
23.9607 | -100.0389
23.9644 | -100.0462
23.9682 | -100.0531
23.9719 | -100.058
23.9757 | -100.0625
23.9794 | -100.0671
23.9832 | -100.0711
23.9869 | -100.0761
23.9906 | -100.0798
23.9944 | -100.0833
23.9981 | -100.085
24.0019 | -100.0881
24.0056 | -100.0917
24.0094 | -100.0949
24.0131 | -100.0996
24.0168 | -100.1054
24.0206 | -100.1102
24.0243 | -100.1129
24.0281 | -100.1154
24.0318 | -100.1175
24.0356 | -100.1201
24.0393 | -100.1243
24.043 | -100.1291
24.0468 | -100.1334
24.0505 | -100.1375
24.0543 | -100.1403
24.058 | -100.1414
24.0618 | -100.1418
24.0655 | -100.1406
24.0692 | -100.1392
24.073 | -100.1391
24.0767 | -100.1386
24.0805 | -100.137
24.0842 | -100.1364
24.088 | -100.1369
24.0917 | -100.1389
24.0954 | -100.1411
24.0992 | -100.1404
24.1029 | -100.1382
24.1067 | -100.1351
24.1104 | -100.1319
24.1142 | -100.1291
24.1179 | -100.127
24.1216 | -100.1249
24.1254 | -100.1224
24.1291 | -100.1188
24.1329 | -100.1147
24.1366 | -100.1102
24.1404 | -100.1052
24.1441 | -100.1019
24.1478 | -100.0995
24.1516 | -100.0956
24.1553 | -100.0914
24.1591 | -100.0882
24.1628 | -100.0846
24.1666 | -100.0835
24.1703 | -100.0862
24.174 | -100.0862
24.1778 | -100.0842
24.1815 | -100.081
24.1853 | -100.0789
24.189 | -100.0803
24.1927 | -100.0827
24.1965 | -100.0848
24.2002 | -100.0862
24.204 | -100.0876
24.2077 | -100.0899
24.2115 | -100.0899
24.2152 | -100.0887
24.2189 | -100.0884
24.2227 | -100.09
24.2264 | -100.0912
24.2302 | -100.0881
24.2339 | -100.0846
24.2377 | -100.0848
24.2414 | -100.0877
24.2451 | -100.0942
24.2489 | -100.0996
24.2526 | -100.1045
24.2564 | -100.1098
24.2601 | -100.1108
24.2639 | -100.1015
24.2676 | -100.0795
24.2713 | -100.0422
24.2751 | -100.0067
24.2788 | -99.9677
24.2826 | -99.9275
24.2863 | -99.8872
24.2901 | -99.8524
24.2938 | -99.8203
24.2975 | -99.7572
24.3013 | -99.7192
24.305 | -99.6791
24.3088 | -99.6373
24.3125 | -99.5929
24.3163 | -99.5463
24.32 | -99.4958
24.3237 | -99.4434
24.3275 | -99.3902
24.3312 | -99.334
24.335 | -99.2744
24.3387 | -99.213
24.3425 | -99.1511
24.3462 | -99.0864
24.3499 | -99.0215
24.3537 | -98.9554
24.3574 | -98.8879
24.3612 | -98.8206
24.3649 | -98.7535
24.3687 | -98.6862
24.3724 | -98.6199
24.3761 | -98.5549
24.3799 | -98.493
24.3836 | -98.4327
24.3874 | -98.374
24.3911 | -98.3169
24.3949 | -98.2638
24.3986 | -98.2144
24.4023 | -98.1684
24.4061 | -98.1248
24.4098 | -98.084
24.4136 | -98.0475
24.4173 | -98.0143
24.4211 | -97.9866
24.4248 | -97.9646
24.4285 | -97.9467
24.4323 | -97.9343
24.436 | -97.9255
24.4398 | -97.9196
24.4435 | -97.9193
24.4473 | -97.9248
24.451 | -97.9346
24.4547 | -97.9489
24.4585 | -97.9675
24.4622 | -97.9904
24.466 | -98.0182
24.4697 | -98.0511
24.4734 | -98.0874
24.4772 | -98.127
24.4809 | -98.1702
24.4847 | -98.2169
24.4884 | -98.2666
24.4922 | -98.3195
24.4959 | -98.3747
24.4996 | -98.4327
24.5034 | -98.4933
24.5071 | -98.5562
24.5109 | -98.6195
24.5146 | -98.6836
24.5184 | -98.7508
24.5221 | -98.8196
24.5258 | -98.8898
24.5296 | -98.9599
24.5333 | -99.0283
24.5371 | -99.0963
24.5408 | -99.1628
24.5446 | -99.2299
24.5483 | -99.2947
24.552 | -99.357
24.5558 | -99.4199
24.5595 | -99.4821
24.5633 | -99.5385
24.567 | -99.5837
24.5708 | -99.6178
24.5745 | -99.648
24.5782 | -99.6799
24.582 | -99.7118
24.5857 | -99.7387
24.5895 | -99.7628
24.5932 | -99.7828
24.597 | -99.8005
24.6007 | -99.8145
24.6044 | -99.8282
24.6082 | -99.8459
24.6119 | -99.8664
24.6157 | -99.8915
24.6194 | -99.9044
24.6232 | -99.9108
24.6269 | -99.9087
24.6306 | -99.8992
24.6344 | -99.8774
24.6381 | -99.8487
24.6419 | -99.8201
24.6456 | -99.7893
24.6494 | -99.7555
24.6531 | -99.7206
24.6568 | -99.6866
24.6606 | -99.6511
24.6643 | -99.6137
24.6681 | -99.5772
24.6718 | -99.5386
24.6756 | -99.4997
24.6793 | -99.4634
24.683 | -99.4287
24.6868 | -99.3942
24.6905 | -99.362
24.6943 | -99.3305
24.698 | -99.3002
24.7018 | -99.2731
24.7055 | -99.2497
24.7092 | -99.2269
24.713 | -99.2067
24.7167 | -99.1891
24.7205 | -99.173
24.7242 | -99.1582
24.728 | -99.1447
24.7317 | -99.1334
24.7354 | -99.1244
24.7392 | -99.1199
24.7429 | -99.1176
24.7467 | -99.118
24.7504 | -99.1217
24.7542 | -99.1287
24.7579 | -99.1379
24.7616 | -99.1494
24.7654 | -99.163
24.7691 | -99.1765
24.7729 | -99.1917
24.7766 | -99.2078
24.7804 | -99.2265
24.7841 | -99.249
24.7878 | -99.2746
24.7916 | -99.3014
24.7953 | -99.3268
24.7991 | -99.3544
24.8028 | -99.3827
24.8066 | -99.4118
24.8103 | -99.4393
24.814 | -99.468
24.8178 | -99.4972
24.8215 | -99.5255
24.8253 | -99.552
24.829 | -99.5767
24.8327 | -99.6004
24.8365 | -99.6102
24.8402 | -99.6306
24.844 | -99.6509
24.8477 | -99.671
24.8515 | -99.6887
24.8552 | -99.7054
24.8589 | -99.7207
24.8627 | -99.7363
24.8664 | -99.7495
24.8702 | -99.7604
24.8739 | -99.7647
24.8777 | -99.7679
24.8814 | -99.7706
24.8851 | -99.7701
24.8889 | -99.7631
24.8926 | -99.7576
24.8964 | -99.7514
24.9001 | -99.7443
24.9039 | -99.736
24.9076 | -99.7238
24.9113 | -99.7117
24.9151 | -99.7015
24.9188 | -99.6968
24.9226 | -99.6969
24.9263 | -99.694
24.9301 | -99.6817
24.9338 | -99.6626
24.9375 | -99.6426
24.9413 | -99.6174
24.945 | -99.5924
24.9488 | -99.5661
24.9525 | -99.5401
24.9563 | -99.515
24.96 | -99.4893
24.9637 | -99.4625
24.9675 | -99.4365
24.9712 | -99.4111
24.975 | -99.3857
24.9787 | -99.3605
24.9825 | -99.3342
24.9862 | -99.308
24.9899 | -99.2838
24.9937 | -99.261
24.9974 | -99.2403
25.0012 | -99.2206
25.0049 | -99.203
25.0087 | -99.1847
25.0124 | -99.1676
25.0161 | -99.154
25.0199 | -99.1425
25.0236 | -99.1332
25.0274 | -99.1247
25.0311 | -99.115
25.0349 | -99.1058
25.0386 | -99.0984
25.0423 | -99.0917
25.0461 | -99.0876
25.0498 | -99.0847
25.0536 | -99.0817
25.0573 | -99.0805
25.0611 | -99.0818
25.0648 | -99.0831
25.0685 | -99.085
25.0723 | -99.0894
25.076 | -99.0972
25.0798 | -99.107
25.0835 | -99.1187
25.0873 | -99.1326
25.091 | -99.1461
25.0947 | -99.1611
25.0985 | -99.1778
25.1022 | -99.1957
25.106 | -99.2142
25.1097 | -99.2349
25.1135 | -99.2569
25.1172 | -99.2804
25.1209 | -99.3058
25.1247 | -99.3327
25.1284 | -99.3616
25.1322 | -99.388
25.1359 | -99.4132
25.1396 | -99.4395
25.1434 | -99.4675
25.1471 | -99.498
25.1509 | -99.5277
25.1546 | -99.5587
25.1584 | -99.5905
25.1621 | -99.6252
25.1658 | -99.6597
25.1696 | -99.6926
25.1733 | -99.7254
25.1771 | -99.7612
25.1808 | -99.7969
25.1846 | -99.8302
25.1883 | -99.8629
25.192 | -99.8966
25.1958 | -99.9292
25.1995 | -99.9586
25.2033 | -99.9871
25.207 | -100.0153
25.2108 | -100.0399
25.2145 | -100.0645
25.2182 | -100.0901
25.222 | -100.1161
25.2257 | -100.1409
25.2295 | -100.1695
25.2332 | -100.1943
25.237 | -100.2162
25.2407 | -100.2391
25.2444 | -100.2658
25.2482 | -100.2888
25.2519 | -100.3085
25.2557 | -100.3277
25.2594 | -100.3454
25.2632 | -100.3629
25.2669 | -100.3786
25.2706 | -100.3924
25.2744 | -100.4022
25.2781 | -100.4117
25.2819 | -100.421
25.2856 | -100.4284
25.2894 | -100.4337
25.2931 | -100.4387
25.2968 | -100.4435
25.3006 | -100.4466
25.3043 | -100.4513
25.3081 | -100.4534
25.3118 | -100.4526
25.3156 | -100.4553
25.3193 | -100.4594
25.323 | -100.4625
25.3268 | -100.4631
25.3305 | -100.4623
25.3343 | -100.461
25.338 | -100.46
25.3418 | -100.459
25.3455 | -100.4563
25.3492 | -100.4511
25.353 | -100.4455
25.3567 | -100.4403
25.3605 | -100.4328
25.3642 | -100.4253
25.368 | -100.42
25.3717 | -100.4143
25.3754 | -100.4082
25.3792 | -100.3987
25.3829 | -100.3895
25.3867 | -100.3813
25.3904 | -100.3733
25.3942 | -100.3639
25.3979 | -100.3564
25.4016 | -100.3504
25.4054 | -100.3426
25.4091 | -100.3355
25.4129 | -100.3289
25.4166 | -100.3191
25.4204 | -100.31
25.4241 | -100.3034
25.4278 | -100.2972
25.4316 | -100.2903
25.4353 | -100.2811
25.4391 | -100.2693
25.4428 | -100.2587
25.4465 | -100.2486
25.4503 | -100.2355
25.454 | -100.2218
25.4578 | -100.2115
25.4615 | -100.2017
25.4653 | -100.1915
25.469 | -100.1818
25.4727 | -100.1723
25.4765 | -100.1612
25.4802 | -100.1487
25.484 | -100.1361
25.4877 | -100.124
25.4915 | -100.1157
25.4952 | -100.1069
25.4989 | -100.0974
25.5027 | -100.0877
25.5064 | -100.0772
25.5102 | -100.0664
25.5139 | -100.0542
25.5177 | -100.0418
25.5214 | -100.0298
25.5251 | -100.0188
25.5289 | -100.0097
25.5326 | -100.0006
25.5364 | -99.9896
25.5401 | -99.9811
25.5439 | -99.9733
25.5476 | -99.9641
25.5513 | -99.9546
25.5551 | -99.9466
25.5588 | -99.9402
25.5626 | -99.9329
25.5663 | -99.9262
25.5701 | -99.9214
25.5738 | -99.916
25.5775 | -99.9103
25.5813 | -99.9035
25.585 | -99.8958
25.5888 | -99.8882
25.5925 | -99.8833
25.5963 | -99.8787
25.6 | -99.8759
25.6037 | -99.8733
25.6075 | -99.8697
25.6112 | -99.8659
25.615 | -99.8627
25.6187 | -99.8613
25.6225 | -99.86
25.6262 | -99.8559
25.6299 | -99.8517
25.6337 | -99.8508
25.6374 | -99.8495
25.6412 | -99.8495
25.6449 | -99.8494
25.6487 | -99.849
25.6524 | -99.8479
25.6561 | -99.8455
25.6599 | -99.8408
25.6636 | -99.8347
25.6674 | -99.8291
25.6711 | -99.8231
25.6749 | -99.8143
25.6786 | -99.8062
25.6823 | -99.7981
25.6861 | -99.7898
25.6898 | -99.7813
25.6936 | -99.7734
25.6973 | -99.7651
25.7011 | -99.756
25.7048 | -99.7453
25.7085 | -99.735
25.7123 | -99.7233
25.716 | -99.7124
25.7198 | -99.7007
25.7235 | -99.6871
25.7273 | -99.6694
25.731 | -99.6509
25.7347 | -99.6295
25.7385 | -99.609
25.7422 | -99.5913
25.746 | -99.5705
25.7497 | -99.5502
25.7535 | -99.5316
25.7572 | -99.5108
25.7609 | -99.4872
25.7647 | -99.464
25.7684 | -99.4398
25.7722 | -99.4143
25.7759 | -99.3905
25.7796 | -99.3658
25.7834 | -99.3407
25.7871 | -99.3168
25.7909 | -99.2933
25.7946 | -99.2688
25.7984 | -99.2443
25.8021 | -99.222
25.8058 | -99.1991
25.8096 | -99.1759
25.8133 | -99.1559
25.8171 | -99.1371
25.8208 | -99.117
25.8246 | -99.095
25.8283 | -99.0719
25.832 | -99.05
25.8358 | -99.0295
25.8395 | -99.0095
25.8433 | -98.99
25.847 | -98.9718
25.8508 | -98.955
25.8545 | -98.9383
25.8582 | -98.924
25.862 | -98.9116
25.8657 | -98.8999
25.8695 | -98.8882
25.8732 | -98.8782
25.877 | -98.8693
25.8807 | -98.8609
25.8844 | -98.8532
25.8882 | -98.8451
25.8919 | -98.8375
25.8957 | -98.8327
25.8994 | -98.8292
25.9032 | -98.8253
25.9069 | -98.8226
25.9106 | -98.8202
25.9144 | -98.8245
25.9181 | -98.822
25.9219 | -98.8205
25.9256 | -98.8196
25.9294 | -98.8207
25.9331 | -98.8214
25.9368 | -98.8214
25.9406 | -98.8211
25.9443 | -98.8215
25.9481 | -98.8235
25.9518 | -98.8264
25.9556 | -98.8307
25.9593 | -98.8344
25.963 | -98.8392
25.9668 | -98.8445
25.9705 | -98.8494
25.9743 | -98.8549
25.978 | -98.8625
25.9818 | -98.8691
25.9855 | -98.8741
25.9892 | -98.8814
25.993 | -98.8913
25.9967 | -98.9017
26.0005 | -98.9121
26.0042 | -98.9226
26.008 | -98.9318
26.0117 | -98.9408
26.0154 | -98.9497
26.0192 | -98.9574
26.0229 | -98.9681
26.0267 | -98.9814
26.0304 | -98.9948
26.0342 | -99.0097
26.0379 | -99.024
26.0416 | -99.0363
26.0454 | -99.0467
26.0491 | -99.0558
26.0529 | -99.0652
26.0566 | -99.0726
26.0604 | -99.0777
26.0641 | -99.0846
26.0678 | -99.0906
26.0716 | -99.0959
26.0753 | -99.1005
26.0791 | -99.1038
26.0828 | -99.1048
26.0865 | -99.1068
26.0903 | -99.1098
26.094 | -99.1126
26.0978 | -99.1136
26.1015 | -99.1157
26.1053 | -99.115
26.109 | -99.1121
26.1127 | -99.1094
26.1165 | -99.1035
26.1202 | -99.0946
26.124 | -99.0867
26.1277 | -99.0749
26.1315 | -99.0612
26.1352 | -99.0503
26.1389 | -99.0385
26.1427 | -99.0239
26.1464 | -99.0114
26.1502 | -98.9985
26.1539 | -98.9827
26.1577 | -98.9683
26.1614 | -98.9544
26.1651 | -98.9409
26.1689 | -98.9266
26.1726 | -98.9121
26.1764 | -98.8992
26.1801 | -98.8858
26.1839 | -98.8719
26.1876 | -98.8575
26.1913 | -98.8456
26.1951 | -98.8362
26.1988 | -98.8253
26.2026 | -98.8134
26.2063 | -98.8041
26.2101 | -98.7944
26.2138 | -98.7829
26.2175 | -98.7736
26.2213 | -98.7644
26.225 | -98.7553
26.2288 | -98.7456
26.2325 | -98.7355
26.2363 | -98.7285
26.24 | -98.7221
26.2437 | -98.7168
26.2475 | -98.7125
26.2512 | -98.7089
26.255 | -98.7064
26.2587 | -98.7038
26.2625 | -98.7016
26.2662 | -98.7004
26.2699 | -98.6997
26.2737 | -98.7003
26.2774 | -98.7023
26.2812 | -98.7036
26.2849 | -98.7055
26.2887 | -98.7076
26.2924 | -98.7095
26.2961 | -98.7126
26.2999 | -98.715
26.3036 | -98.7205
26.3074 | -98.7282
26.3111 | -98.7365
26.3149 | -98.7447
26.3186 | -98.7525
26.3223 | -98.7601
26.3261 | -98.7661
26.3298 | -98.7709
26.3336 | -98.7743
26.3373 | -98.7753
26.3411 | -98.7787
26.3448 | -98.7853
26.3485 | -98.7909
26.3523 | -98.7971
26.356 | -98.8038
26.3598 | -98.8088
26.3635 | -98.8134
26.3673 | -98.8174
26.371 | -98.8193
26.3747 | -98.821
26.3785 | -98.8244
26.3822 | -98.8286
26.386 | -98.8327
26.3897 | -98.8374
26.3935 | -98.8406
26.3972 | -98.8425
26.4009 | -98.8438
26.4047 | -98.8433
26.4084 | -98.8428
26.4122 | -98.843
26.4159 | -98.8421
26.4196 | -98.8405
26.4234 | -98.8379
26.4271 | -98.8355
26.4309 | -98.8336
26.4346 | -98.8288
26.4384 | -98.8238
26.4421 | -98.8175
26.4458 | -98.8102
26.4496 | -98.8024
26.4533 | -98.7962
26.4571 | -98.7878
26.4608 | -98.7809
26.4646 | -98.7721
26.4683 | -98.7652
26.472 | -98.7583
26.4758 | -98.7492
26.4795 | -98.7407
26.4833 | -98.7331
26.487 | -98.7249
26.4908 | -98.7189
26.4945 | -98.7139
26.4982 | -98.71
26.502 | -98.7085
26.5057 | -98.7061
26.5095 | -98.7027
26.5132 | -98.6993
26.517 | -98.6966
26.5207 | -98.6951
26.5244 | -98.6943
26.5282 | -98.6963
26.5319 | -98.6972
26.5357 | -98.698
26.5394 | -98.7002
26.5432 | -98.7041
26.5469 | -98.7101
26.5506 | -98.7162
26.5544 | -98.7235
26.5581 | -98.7327
26.5619 | -98.7432
26.5656 | -98.7548
26.5694 | -98.7679
26.5731 | -98.7836
26.5768 | -98.8008
26.5806 | -98.8181
26.5843 | -98.8369
26.5881 | -98.8573
26.5918 | -98.8758
26.5956 | -98.896
26.5993 | -98.9169
26.603 | -98.9388
26.6068 | -98.9618
26.6105 | -98.9848
26.6143 | -99.0083
26.618 | -99.0328
26.6218 | -99.0563
26.6255 | -99.0794
26.6292 | -99.1031
26.633 | -99.1284
26.6367 | -99.1529
26.6405 | -99.176
26.6442 | -99.198
26.648 | -99.218
26.6517 | -99.2392
26.6554 | -99.2621
26.6592 | -99.2849
26.6629 | -99.3093
26.6667 | -99.3304
26.6704 | -99.3498
26.6742 | -99.3712
26.6779 | -99.3906
26.6816 | -99.4086
26.6854 | -99.4257
26.6891 | -99.4395
26.6929 | -99.4542
26.6966 | -99.4703
26.7004 | -99.4844
26.7041 | -99.4988
26.7078 | -99.5136
26.7116 | -99.5279
26.7153 | -99.5375
26.7191 | -99.5443
26.7228 | -99.5493
26.7265 | -99.5526
26.7303 | -99.5543
26.734 | -99.5526
26.7378 | -99.5493
26.7415 | -99.5455
26.7453 | -99.5402
26.749 | -99.5317
26.7527 | -99.522
26.7565 | -99.5118
26.7602 | -99.5007
26.764 | -99.4872
26.7677 | -99.4725
26.7715 | -99.4554
26.7752 | -99.4354
26.7789 | -99.416
26.7827 | -99.3969
26.7864 | -99.3783
26.7902 | -99.3596
26.7939 | -99.3398
26.7977 | -99.3206
26.8014 | -99.3005
26.8051 | -99.2793
26.8089 | -99.2595
26.8126 | -99.2381
26.8164 | -99.2175
26.8201 | -99.1974
26.8239 | -99.1787
26.8276 | -99.1584
26.8313 | -99.1364
26.8351 | -99.1124
26.8388 | -99.0902
26.8426 | -99.07
26.8463 | -99.048
26.8501 | -99.026
26.8538 | -99.0064
26.8575 | -98.9865
26.8613 | -98.9645
26.865 | -98.9426
26.8688 | -98.9213
26.8725 | -98.9018
26.8763 | -98.8843
26.88 | -98.8689
26.8837 | -98.8532
26.8875 | -98.8377
26.8912 | -98.8235
26.895 | -98.8115
26.8987 | -98.7988
26.9025 | -98.786
26.9062 | -98.7736
26.9099 | -98.7627
26.9137 | -98.7553
26.9174 | -98.7485
26.9212 | -98.7417
26.9249 | -98.7364
26.9287 | -98.7331
26.9324 | -98.7307
26.9361 | -98.7291
26.9399 | -98.7275
26.9436 | -98.7268
26.9474 | -98.7284
26.9511 | -98.7309
26.9549 | -98.7336
26.9586 | -98.7388
26.9623 | -98.7458
26.9661 | -98.7552
26.9698 | -98.7647
26.9736 | -98.7725
26.9773 | -98.7815
26.9811 | -98.7919
26.9848 | -98.8027
26.9885 | -98.8152
26.9923 | -98.8267
26.996 | -98.8436
26.9998 | -98.8562
27.0035 | -98.8703
27.0073 | -98.8851
27.011 | -98.8985
27.0147 | -98.9118
27.0185 | -98.9258
27.0222 | -98.9399
27.026 | -98.9548
27.0297 | -98.9706
27.0334 | -98.984
27.0372 | -98.9948
27.0409 | -99.0036
27.0447 | -99.0119
27.0484 | -99.0229
27.0522 | -99.0327
27.0559 | -99.0385
27.0596 | -99.0441
27.0634 | -99.0504
27.0671 | -99.0561
27.0709 | -99.0627
27.0746 | -99.0659
27.0784 | -99.0679
27.0821 | -99.0718
27.0858 | -99.0738
27.0896 | -99.075
27.0933 | -99.0754
27.0971 | -99.0759
27.1008 | -99.0751
27.1046 | -99.0741
27.1083 | -99.074
27.112 | -99.0743
27.1158 | -99.0731
27.1195 | -99.0713
27.1233 | -99.0686
27.127 | -99.0654
27.1308 | -99.0619
27.1345 | -99.056
27.1382 | -99.0485
27.142 | -99.04
27.1457 | -99.0322
27.1495 | -99.0255
27.1532 | -99.0202
27.157 | -99.0153
27.1607 | -99.0089
27.1644 | -99.002
27.1682 | -98.9938
27.1719 | -98.9878
27.1757 | -98.9829
27.1794 | -98.9751
27.1832 | -98.9681
27.1869 | -98.9634
27.1906 | -98.9566
27.1944 | -98.9517
27.1981 | -98.9481
27.2019 | -98.9437
27.2056 | -98.9383
27.2094 | -98.9327
27.2131 | -98.928
27.2168 | -98.9253
27.2206 | -98.9229
27.2243 | -98.9212
27.2281 | -98.9221
27.2318 | -98.9235
27.2356 | -98.9239
27.2393 | -98.9235
27.243 | -98.9221
27.2468 | -98.9193
27.2505 | -98.917
27.2543 | -98.9133
27.258 | -98.9099
27.2618 | -98.9054
27.2655 | -98.9025
27.2692 | -98.9008
27.273 | -98.8987
27.2767 | -98.8954
27.2805 | -98.8908
27.2842 | -98.8871
27.288 | -98.8834
27.2917 | -98.8781
27.2954 | -98.8703
27.2992 | -98.8631
27.3029 | -98.8558
27.3067 | -98.8458
27.3104 | -98.8335
27.3142 | -98.8191
27.3179 | -98.8024
27.3216 | -98.7885
27.3254 | -98.775
27.3291 | -98.7594
27.3329 | -98.7408
27.3366 | -98.721
27.3404 | -98.7029
27.3441 | -98.6856
27.3478 | -98.6645
27.3516 | -98.6447
27.3553 | -98.6277
27.3591 | -98.61
27.3628 | -98.5906
27.3666 | -98.5698
27.3703 | -98.5487
27.374 | -98.5279
27.3778 | -98.5067
27.3815 | -98.4826
27.3853 | -98.4574
27.389 | -98.4341
27.3927 | -98.4112
27.3965 | -98.3878
27.4002 | -98.3647
27.404 | -98.3427
27.4077 | -98.3209
27.4115 | -98.3001
27.4152 | -98.2768
27.4189 | -98.2537
27.4227 | -98.2339
27.4264 | -98.2155
27.4302 | -98.1954
27.4339 | -98.174
27.4377 | -98.1546
27.4414 | -98.1367
27.4451 | -98.1182
27.4489 | -98.0987
27.4526 | -98.0802
27.4564 | -98.0641
27.4601 | -98.05
27.4639 | -98.0356
27.4676 | -98.0205
27.4713 | -98.0066
27.4751 | -97.9933
27.4788 | -97.9818
27.4826 | -97.9715
27.4863 | -97.9607
27.4901 | -97.9485
27.4938 | -97.9374
27.4975 | -97.9256
27.5013 | -97.9151
27.505 | -97.9064
27.5088 | -97.8981
27.5125 | -97.8906
27.5163 | -97.8848
27.52 | -97.8812
27.5237 | -97.8784
27.5275 | -97.8767
27.5312 | -97.8761
27.535 | -97.8796
27.5387 | -97.8764
27.5425 | -97.8723
27.5462 | -97.87
27.5499 | -97.8684
27.5537 | -97.8673
27.5574 | -97.8671
27.5612 | -97.8668
27.5649 | -97.8671
27.5687 | -97.8688
27.5724 | -97.8711
27.5761 | -97.8717
27.5799 | -97.8739
27.5836 | -97.876
27.5874 | -97.8765
27.5911 | -97.8782
27.5949 | -97.8799
27.5986 | -97.8835
27.6023 | -97.8887
27.6061 | -97.8932
27.6098 | -97.8975
27.6136 | -97.9007
27.6173 | -97.9022
27.6211 | -97.9044
27.6248 | -97.9066
27.6285 | -97.9093
27.6323 | -97.9132
27.636 | -97.9189
27.6398 | -97.9227
27.6435 | -97.9252
27.6473 | -97.9265
27.651 | -97.9279
27.6547 | -97.9296
27.6585 | -97.9302
27.6622 | -97.9308
27.666 | -97.9318
27.6697 | -97.9324
27.6734 | -97.9308
27.6772 | -97.9282
27.6809 | -97.926
27.6847 | -97.9246
27.6884 | -97.9226
27.6922 | -97.9189
27.6959 | -97.9138
27.6996 | -97.9119
27.7034 | -97.913
27.7071 | -97.9131
27.7109 | -97.9102
27.7146 | -97.9066
27.7184 | -97.902
27.7221 | -97.8993
27.7258 | -97.8976
27.7296 | -97.8933
27.7333 | -97.8882
27.7371 | -97.8844
27.7408 | -97.8791
27.7446 | -97.872
27.7483 | -97.8662
27.752 | -97.8617
27.7558 | -97.8572
27.7595 | -97.8522
27.7633 | -97.8466
27.767 | -97.8392
27.7708 | -97.8345
27.7745 | -97.8311
27.7782 | -97.8233
27.782 | -97.8166
27.7857 | -97.813
27.7895 | -97.8074
27.7932 | -97.8003
27.797 | -97.7946
27.8007 | -97.7865
27.8044 | -97.7805
27.8082 | -97.7757
27.8119 | -97.7683
27.8157 | -97.7589
27.8194 | -97.751
27.8232 | -97.7439
27.8269 | -97.7349
27.8306 | -97.7257
27.8344 | -97.7184
27.8381 | -97.7122
27.8419 | -97.7061
27.8456 | -97.6968
27.8494 | -97.6874
27.8531 | -97.6768
27.8568 | -97.664
27.8606 | -97.6482
27.8643 | -97.6341
27.8681 | -97.6204
27.8718 | -97.6076
27.8756 | -97.5976
27.8793 | -97.5869
27.883 | -97.5757
27.8868 | -97.5668
27.8905 | -97.5584
27.8943 | -97.5477
27.898 | -97.5369
27.9018 | -97.5246
27.9055 | -97.5121
27.9092 | -97.4979
27.913 | -97.4834
27.9167 | -97.4693
27.9205 | -97.4537
27.9242 | -97.4387
27.928 | -97.4168
27.9317 | -97.3861
27.9354 | -97.3655
27.9392 | -97.3491
27.9429 | -97.3307
27.9467 | -97.3112
27.9504 | -97.2922
27.9542 | -97.2705
27.9579 | -97.2488
27.9616 | -97.2296
27.9654 | -97.2089
27.9691 | -97.1893
27.9729 | -97.1702
27.9766 | -97.1497
27.9804 | -97.1298
27.9841 | -97.1092
27.9878 | -97.0888
27.9916 | -97.0682
27.9953 | -97.0477
27.9991 | -97.0239
28.0028 | -96.9977