## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 267 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 259
Hood | 02-92B/P | 186
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1012.3 | hPa |  | PRC PRO 01
Humidity | 56 | %/Hr
Internal temperature of calibration bench | 23.68 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1021 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.2 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.02 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3062 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 12/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 18/07/2018
Technical Manager | Anthony HUE | Approval |  | 24/07/2018
CEO | Robert GLEMOT | Approval |  | 24/07/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4703 | 57.6495
0.012209 | -36.0285 | 38.4662
0.018313 | -34.9562 | 28.3178
0.024417 | -34.4889 | 22.0615
0.030521 | -34.3771 | 18.2884
0.036626 | -34.2317 | 15.2266
0.04273 | -34.1352 | 12.8528
0.048834 | -34.0882 | 11.3274
0.054938 | -34.0219 | 10.1848
0.061043 | -33.9596 | 9.0962
0.067147 | -33.9491 | 8.273
0.073251 | -33.9501 | 7.5123
0.079355 | -33.9687 | 6.8503
0.08546 | -33.9826 | 6.2545
0.091564 | -33.9775 | 5.8131
0.097668 | -33.959 | 5.5656
0.10377 | -33.9364 | 5.3182
0.10988 | -33.9187 | 5.0407
0.11598 | -33.9093 | 4.7471
0.12209 | -33.911 | 4.462
0.12819 | -33.9211 | 4.1768
0.13429 | -33.9322 | 3.8871
0.1404 | -33.9354 | 3.6076
0.1465 | -33.9373 | 3.3488
0.15261 | -33.9344 | 3.1279
0.15871 | -33.9276 | 2.9654
0.16482 | -33.9192 | 2.8628
0.17092 | -33.9167 | 2.8241
0.17702 | -33.9144 | 2.7519
0.18313 | -33.9133 | 2.6738
0.18923 | -33.9128 | 2.5646
0.19534 | -33.9113 | 2.4374
0.20144 | -33.9099 | 2.2947
0.20754 | -33.907 | 2.1737
0.21365 | -33.9059 | 2.059
0.21975 | -33.9054 | 1.9612
0.22586 | -33.9065 | 1.8741
0.23196 | -33.9089 | 1.802
0.23807 | -33.913 | 1.7398
0.24417 | -33.917 | 1.6814
0.25027 | -33.9209 | 1.6193
0.25638 | -33.9233 | 1.5576
0.26248 | -33.925 | 1.491
0.26859 | -33.9259 | 1.4257
0.27469 | -33.9262 | 1.3545
0.2808 | -33.9272 | 1.2805
0.2869 | -33.928 | 1.2093
0.293 | -33.9289 | 1.1419
0.29911 | -33.9298 | 1.0817
0.30521 | -33.9307 | 1.0223
0.31132 | -33.9314 | 0.98341
0.31742 | -33.9326 | 0.94038
0.32353 | -33.9328 | 0.90524
0.32963 | -33.9337 | 0.87607
0.33573 | -33.9346 | 0.85568
0.34184 | -33.9359 | 0.83424
0.34794 | -33.938 | 0.81048
0.35405 | -33.9401 | 0.78643
0.36015 | -33.9426 | 0.76115
0.36626 | -33.9446 | 0.73375
0.37236 | -33.9467 | 0.69705
0.37846 | -33.9488 | 0.65279
0.38457 | -33.9497 | 0.59554
0.39067 | -33.9508 | 0.54886
0.39678 | -33.9514 | 0.49687
0.40288 | -33.9519 | 0.44263
0.40899 | -33.9518 | 0.39058
0.41509 | -33.9511 | 0.33952
0.42119 | -33.951 | 0.29603
0.4273 | -33.951 | 0.24959
0.4334 | -33.9509 | 0.20321
0.43951 | -33.9511 | 0.15956
0.44561 | -33.952 | 0.11951
0.45172 | -33.9535 | 0.087607
0.45782 | -33.9543 | 0.050271
0.46392 | -33.9555 | 0.013829
0.47003 | -33.9565 | -0.018785
0.47613 | -33.9573 | -0.053184
0.48224 | -33.9578 | -0.088282
0.48834 | -33.9568 | -0.12341
0.49445 | -33.9565 | -0.15904
0.50055 | -33.9559 | -0.19144
0.50665 | -33.9551 | -0.226
0.51276 | -33.954 | -0.26041
0.51886 | -33.9529 | -0.29471
0.52497 | -33.9516 | -0.31908
0.53107 | -33.9508 | -0.34882
0.53717 | -33.9501 | -0.37719
0.54328 | -33.9492 | -0.40023
0.54938 | -33.949 | -0.4228
0.55549 | -33.949 | -0.44384
0.56159 | -33.9501 | -0.45541
0.5677 | -33.9504 | -0.46999
0.5738 | -33.9514 | -0.48635
0.5799 | -33.9522 | -0.50132
0.58601 | -33.9533 | -0.51138
0.59211 | -33.955 | -0.51767
0.59822 | -33.9564 | -0.52712
0.60432 | -33.9579 | -0.53923
0.61043 | -33.9592 | -0.54713
0.61653 | -33.9604 | -0.55742
0.62263 | -33.962 | -0.56671
0.62874 | -33.9626 | -0.57722
0.63484 | -33.9634 | -0.58639
0.64095 | -33.9639 | -0.59646
0.64705 | -33.9643 | -0.60606
0.65316 | -33.9644 | -0.61842
0.65926 | -33.9645 | -0.63417
0.66536 | -33.9633 | -0.65156
0.67147 | -33.9629 | -0.66462
0.67757 | -33.9625 | -0.67882
0.68368 | -33.9617 | -0.69771
0.68978 | -33.9609 | -0.71477
0.69589 | -33.9599 | -0.73377
0.70199 | -33.9589 | -0.75663
0.70809 | -33.958 | -0.77581
0.7142 | -33.9573 | -0.79636
0.7203 | -33.9562 | -0.81561
0.72641 | -33.9555 | -0.83826
0.73251 | -33.9548 | -0.85841
0.73862 | -33.955 | -0.88039
0.74472 | -33.9546 | -0.90055
0.75082 | -33.9544 | -0.92287
0.75693 | -33.9547 | -0.9449
0.76303 | -33.9552 | -0.96705
0.76914 | -33.957 | -0.98774
0.77524 | -33.958 | -1.0075
0.78135 | -33.9593 | -1.0278
0.78745 | -33.9606 | -1.0478
0.79355 | -33.9619 | -1.0678
0.79966 | -33.9636 | -1.0874
0.80576 | -33.9659 | -1.1071
0.81187 | -33.9675 | -1.1289
0.81797 | -33.9695 | -1.1503
0.82408 | -33.9711 | -1.171
0.83018 | -33.9729 | -1.1917
0.83628 | -33.9748 | -1.2103
0.84239 | -33.976 | -1.2308
0.84849 | -33.9774 | -1.2526
0.8546 | -33.9788 | -1.2725
0.8607 | -33.9801 | -1.2933
0.86681 | -33.9811 | -1.3124
0.87291 | -33.9821 | -1.3312
0.87901 | -33.9826 | -1.3492
0.88512 | -33.9833 | -1.3674
0.89122 | -33.984 | -1.3857
0.89733 | -33.9843 | -1.4013
0.90343 | -33.9848 | -1.419
0.90953 | -33.9849 | -1.4332
0.91564 | -33.9849 | -1.449
0.92174 | -33.9851 | -1.464
0.92785 | -33.9851 | -1.4785
0.93395 | -33.9852 | -1.4923
0.94006 | -33.9853 | -1.5051
0.94616 | -33.985 | -1.5169
0.95226 | -33.9851 | -1.5295
0.95837 | -33.9852 | -1.5433
0.96447 | -33.9854 | -1.5561
0.97058 | -33.9852 | -1.5694
0.97668 | -33.9849 | -1.5826
0.98279 | -33.9846 | -1.5953
0.98889 | -33.9844 | -1.6059
0.99499 | -33.9844 | -1.6193
1.0011 | -33.9842 | -1.6299
1.0072 | -33.9838 | -1.6414
1.0133 | -33.9835 | -1.6521
1.0194 | -33.983 | -1.665
1.0255 | -33.9827 | -1.6786
1.0316 | -33.9821 | -1.6904
1.0377 | -33.9817 | -1.7041
1.0438 | -33.9811 | -1.7167
1.0499 | -33.9803 | -1.732
1.056 | -33.9796 | -1.7462
1.0621 | -33.9789 | -1.7605
1.0682 | -33.9784 | -1.7729
1.0743 | -33.9778 | -1.7883
1.0805 | -33.9772 | -1.8027
1.0866 | -33.9763 | -1.8178
1.0927 | -33.9754 | -1.8322
1.0988 | -33.9748 | -1.8471
1.1049 | -33.974 | -1.8612
1.111 | -33.9732 | -1.8766
1.1171 | -33.9724 | -1.8912
1.1232 | -33.9719 | -1.9064
1.1293 | -33.9715 | -1.9204
1.1354 | -33.9712 | -1.9355
1.1415 | -33.9708 | -1.951
1.1476 | -33.9702 | -1.9665
1.1537 | -33.9697 | -1.9812
1.1598 | -33.9697 | -1.996
1.1659 | -33.9695 | -2.0094
1.172 | -33.9692 | -2.0233
1.1781 | -33.969 | -2.0373
1.1842 | -33.9688 | -2.0515
1.1903 | -33.9691 | -2.0624
1.1964 | -33.969 | -2.0758
1.2025 | -33.9691 | -2.0898
1.2086 | -33.9694 | -2.1023
1.2147 | -33.9694 | -2.1148
1.2209 | -33.9698 | -2.1261
1.227 | -33.9705 | -2.1369
1.2331 | -33.9709 | -2.1489
1.2392 | -33.9712 | -2.1607
1.2453 | -33.9717 | -2.1717
1.2514 | -33.9721 | -2.1816
1.2575 | -33.9726 | -2.1925
1.2636 | -33.9735 | -2.2013
1.2697 | -33.9741 | -2.2104
1.2758 | -33.9747 | -2.2215
1.2819 | -33.9754 | -2.2313
1.288 | -33.9761 | -2.242
1.2941 | -33.9769 | -2.2503
1.3002 | -33.9776 | -2.2582
1.3063 | -33.9781 | -2.2671
1.3124 | -33.9788 | -2.2759
1.3185 | -33.9795 | -2.2848
1.3246 | -33.9802 | -2.2937
1.3307 | -33.981 | -2.3021
1.3368 | -33.9818 | -2.3106
1.3429 | -33.9825 | -2.3193
1.349 | -33.983 | -2.3284
1.3551 | -33.9837 | -2.3375
1.3613 | -33.9842 | -2.3455
1.3674 | -33.9848 | -2.3553
1.3735 | -33.9853 | -2.3649
1.3796 | -33.9858 | -2.3737
1.3857 | -33.9861 | -2.3827
1.3918 | -33.9866 | -2.3927
1.3979 | -33.9872 | -2.4017
1.404 | -33.9877 | -2.4128
1.4101 | -33.9881 | -2.423
1.4162 | -33.9885 | -2.4334
1.4223 | -33.989 | -2.4446
1.4284 | -33.9895 | -2.4545
1.4345 | -33.9897 | -2.4658
1.4406 | -33.99 | -2.4793
1.4467 | -33.9905 | -2.4908
1.4528 | -33.9909 | -2.502
1.4589 | -33.9914 | -2.5137
1.465 | -33.9917 | -2.5245
1.4711 | -33.9922 | -2.5358
1.4772 | -33.9927 | -2.549
1.4833 | -33.9932 | -2.5619
1.4894 | -33.9937 | -2.574
1.4955 | -33.9942 | -2.5859
1.5016 | -33.9948 | -2.5983
1.5078 | -33.9952 | -2.6122
1.5139 | -33.9956 | -2.6253
1.52 | -33.996 | -2.6378
1.5261 | -33.9963 | -2.6504
1.5322 | -33.9967 | -2.6644
1.5383 | -33.9971 | -2.677
1.5444 | -33.9976 | -2.6896
1.5505 | -33.9981 | -2.7031
1.5566 | -33.9986 | -2.7161
1.5627 | -33.9991 | -2.7299
1.5688 | -33.9997 | -2.7435
1.5749 | -34 | -2.7562
1.581 | -34.0003 | -2.7696
1.5871 | -34.0007 | -2.7834
1.5932 | -34.0011 | -2.7963
1.5993 | -34.0015 | -2.8105
1.6054 | -34.0018 | -2.8238
1.6115 | -34.0023 | -2.838
1.6176 | -34.0025 | -2.8506
1.6237 | -34.0028 | -2.8643
1.6298 | -34.003 | -2.877
1.6359 | -34.0034 | -2.8898
1.642 | -34.0038 | -2.9019
1.6482 | -34.004 | -2.915
1.6543 | -34.0043 | -2.9288
1.6604 | -34.0047 | -2.9409
1.6665 | -34.0047 | -2.9535
1.6726 | -34.0049 | -2.9661
1.6787 | -34.0053 | -2.9788
1.6848 | -34.0052 | -2.9916
1.6909 | -34.0054 | -3.0044
1.697 | -34.0057 | -3.0168
1.7031 | -34.0057 | -3.0295
1.7092 | -34.0057 | -3.0419
1.7153 | -34.0058 | -3.0539
1.7214 | -34.0056 | -3.0663
1.7275 | -34.0056 | -3.079
1.7336 | -34.0055 | -3.092
1.7397 | -34.0056 | -3.1045
1.7458 | -34.0056 | -3.1168
1.7519 | -34.0056 | -3.1291
1.758 | -34.0057 | -3.1412
1.7641 | -34.0058 | -3.1529
1.7702 | -34.0058 | -3.166
1.7763 | -34.0057 | -3.1777
1.7824 | -34.0057 | -3.1902
1.7885 | -34.0057 | -3.2028
1.7947 | -34.0058 | -3.2149
1.8008 | -34.0058 | -3.2263
1.8069 | -34.0058 | -3.238
1.813 | -34.0059 | -3.2503
1.8191 | -34.0059 | -3.2631
1.8252 | -34.006 | -3.2759
1.8313 | -34.0061 | -3.2882
1.8374 | -34.0063 | -3.3002
1.8435 | -34.0063 | -3.3122
1.8496 | -34.0064 | -3.3253
1.8557 | -34.0066 | -3.3381
1.8618 | -34.0069 | -3.3503
1.8679 | -34.007 | -3.3624
1.874 | -34.0071 | -3.3744
1.8801 | -34.0073 | -3.3868
1.8862 | -34.0075 | -3.3988
1.8923 | -34.0077 | -3.4109
1.8984 | -34.0078 | -3.4229
1.9045 | -34.0079 | -3.4345
1.9106 | -34.0079 | -3.4466
1.9167 | -34.008 | -3.4588
1.9228 | -34.0081 | -3.4706
1.9289 | -34.0083 | -3.4824
1.9351 | -34.0084 | -3.4931
1.9412 | -34.0087 | -3.5043
1.9473 | -34.009 | -3.5163
1.9534 | -34.0092 | -3.5277
1.9595 | -34.0094 | -3.5394
1.9656 | -34.0096 | -3.5516
1.9717 | -34.0099 | -3.5632
1.9778 | -34.0102 | -3.5745
1.9839 | -34.0104 | -3.5859
1.99 | -34.0106 | -3.5968
1.9961 | -34.0107 | -3.6089
2.0022 | -34.011 | -3.6201
2.0083 | -34.011 | -3.6313
2.0144 | -34.0113 | -3.6426
2.0205 | -34.0116 | -3.6542
2.0266 | -34.0119 | -3.6655
2.0327 | -34.0122 | -3.6769
2.0388 | -34.0126 | -3.6883
2.0449 | -34.0127 | -3.6996
2.051 | -34.013 | -3.7108
2.0571 | -34.0133 | -3.722
2.0632 | -34.0135 | -3.7334
2.0693 | -34.0137 | -3.7454
2.0754 | -34.014 | -3.7567
2.0816 | -34.0143 | -3.7687
2.0877 | -34.0144 | -3.7801
2.0938 | -34.0148 | -3.7921
2.0999 | -34.015 | -3.8037
2.106 | -34.0151 | -3.8152
2.1121 | -34.0153 | -3.8271
2.1182 | -34.0155 | -3.8377
2.1243 | -34.0158 | -3.8484
2.1304 | -34.016 | -3.8594
2.1365 | -34.0163 | -3.871
2.1426 | -34.0164 | -3.8829
2.1487 | -34.0167 | -3.894
2.1548 | -34.0171 | -3.9057
2.1609 | -34.0175 | -3.917
2.167 | -34.0179 | -3.9281
2.1731 | -34.0182 | -3.94
2.1792 | -34.0184 | -3.9509
2.1853 | -34.0188 | -3.9628
2.1914 | -34.019 | -3.974
2.1975 | -34.0194 | -3.9854
2.2036 | -34.0196 | -3.9974
2.2097 | -34.0198 | -4.0087
2.2158 | -34.0202 | -4.0192
2.222 | -34.0204 | -4.0295
2.2281 | -34.0206 | -4.0408
2.2342 | -34.0209 | -4.052
2.2403 | -34.0211 | -4.0625
2.2464 | -34.0213 | -4.0732
2.2525 | -34.0216 | -4.0845
2.2586 | -34.0219 | -4.095
2.2647 | -34.0222 | -4.1053
2.2708 | -34.0224 | -4.1164
2.2769 | -34.0228 | -4.127
2.283 | -34.0229 | -4.138
2.2891 | -34.0231 | -4.1488
2.2952 | -34.0233 | -4.1593
2.3013 | -34.0235 | -4.1701
2.3074 | -34.0237 | -4.1804
2.3135 | -34.0239 | -4.1917
2.3196 | -34.0241 | -4.2017
2.3257 | -34.0243 | -4.2123
2.3318 | -34.0244 | -4.2232
2.3379 | -34.0246 | -4.2339
2.344 | -34.0247 | -4.2451
2.3501 | -34.0249 | -4.2557
2.3562 | -34.025 | -4.2661
2.3623 | -34.0251 | -4.2764
2.3685 | -34.0252 | -4.2868
2.3746 | -34.0253 | -4.298
2.3807 | -34.0255 | -4.3083
2.3868 | -34.0257 | -4.319
2.3929 | -34.0259 | -4.3296
2.399 | -34.0259 | -4.3399
2.4051 | -34.0262 | -4.3504
2.4112 | -34.0263 | -4.3609
2.4173 | -34.0265 | -4.3715
2.4234 | -34.0266 | -4.3822
2.4295 | -34.0267 | -4.393
2.4356 | -34.0269 | -4.4037
2.4417 | -34.027 | -4.4147
2.4478 | -34.0271 | -4.4255
2.4539 | -34.0273 | -4.4357
2.46 | -34.0273 | -4.4471
2.4661 | -34.0275 | -4.4582
2.4722 | -34.0276 | -4.4683
2.4783 | -34.0277 | -4.4786
2.4844 | -34.0278 | -4.4896
2.4905 | -34.0279 | -4.5005
2.4966 | -34.0281 | -4.5112
2.5027 | -34.0281 | -4.5216
2.5089 | -34.0283 | -4.5324
2.515 | -34.0283 | -4.5432
2.5211 | -34.0284 | -4.5539
2.5272 | -34.0284 | -4.5648
2.5333 | -34.0286 | -4.5757
2.5394 | -34.0288 | -4.5864
2.5455 | -34.0289 | -4.5977
2.5516 | -34.029 | -4.608
2.5577 | -34.0292 | -4.6194
2.5638 | -34.0293 | -4.6311
2.5699 | -34.0294 | -4.6425
2.576 | -34.0294 | -4.6541
2.5821 | -34.0295 | -4.6653
2.5882 | -34.0297 | -4.6764
2.5943 | -34.0298 | -4.6872
2.6004 | -34.03 | -4.6981
2.6065 | -34.0301 | -4.7085
2.6126 | -34.0302 | -4.7189
2.6187 | -34.0303 | -4.7294
2.6248 | -34.0305 | -4.7402
2.6309 | -34.0306 | -4.7516
2.637 | -34.0309 | -4.7632
2.6431 | -34.031 | -4.7737
2.6492 | -34.0311 | -4.7851
2.6554 | -34.0313 | -4.7959
2.6615 | -34.0314 | -4.8069
2.6676 | -34.0317 | -4.8183
2.6737 | -34.032 | -4.8294
2.6798 | -34.0323 | -4.8406
2.6859 | -34.0324 | -4.8519
2.692 | -34.0326 | -4.8627
2.6981 | -34.0328 | -4.8738
2.7042 | -34.033 | -4.8849
2.7103 | -34.0332 | -4.8956
2.7164 | -34.0335 | -4.9065
2.7225 | -34.0338 | -4.9179
2.7286 | -34.034 | -4.9289
2.7347 | -34.0342 | -4.9394
2.7408 | -34.0345 | -4.9495
2.7469 | -34.0347 | -4.9607
2.753 | -34.035 | -4.9713
2.7591 | -34.0353 | -4.9826
2.7652 | -34.0355 | -4.9941
2.7713 | -34.0357 | -5.0049
2.7774 | -34.0359 | -5.0158
2.7835 | -34.0361 | -5.0268
2.7896 | -34.0363 | -5.0373
2.7958 | -34.0366 | -5.048
2.8019 | -34.0369 | -5.0584
2.808 | -34.0372 | -5.0692
2.8141 | -34.0375 | -5.0804
2.8202 | -34.0378 | -5.0912
2.8263 | -34.0381 | -5.1021
2.8324 | -34.0384 | -5.1135
2.8385 | -34.0388 | -5.1243
2.8446 | -34.0392 | -5.1351
2.8507 | -34.0396 | -5.1459
2.8568 | -34.0399 | -5.1563
2.8629 | -34.0402 | -5.1669
2.869 | -34.0405 | -5.178
2.8751 | -34.0408 | -5.1887
2.8812 | -34.0412 | -5.1999
2.8873 | -34.0415 | -5.2108
2.8934 | -34.0418 | -5.2218
2.8995 | -34.0422 | -5.2329
2.9056 | -34.0424 | -5.244
2.9117 | -34.0428 | -5.255
2.9178 | -34.0431 | -5.2657
2.9239 | -34.0435 | -5.2773
2.93 | -34.0439 | -5.2887
2.9361 | -34.0442 | -5.2994
2.9423 | -34.0445 | -5.3101
2.9484 | -34.045 | -5.3211
2.9545 | -34.0453 | -5.332
2.9606 | -34.0457 | -5.3434
2.9667 | -34.0461 | -5.3545
2.9728 | -34.0464 | -5.3657
2.9789 | -34.0468 | -5.3768
2.985 | -34.0473 | -5.388
2.9911 | -34.0476 | -5.3992
2.9972 | -34.0481 | -5.4102
3.0033 | -34.0485 | -5.4216
3.0094 | -34.0488 | -5.4327
3.0155 | -34.0493 | -5.4435
3.0216 | -34.0496 | -5.4547
3.0277 | -34.0499 | -5.4663
3.0338 | -34.0503 | -5.4776
3.0399 | -34.0507 | -5.4889
3.046 | -34.051 | -5.5004
3.0521 | -34.0514 | -5.5114
3.0582 | -34.0517 | -5.5233
3.0643 | -34.052 | -5.5347
3.0704 | -34.0523 | -5.5464
3.0765 | -34.0526 | -5.5574
3.0827 | -34.0529 | -5.5686
3.0888 | -34.0532 | -5.5797
3.0949 | -34.0536 | -5.591
3.101 | -34.0539 | -5.6023
3.1071 | -34.0542 | -5.614
3.1132 | -34.0546 | -5.6256
3.1193 | -34.055 | -5.637
3.1254 | -34.0553 | -5.6482
3.1315 | -34.0557 | -5.6601
3.1376 | -34.056 | -5.6714
3.1437 | -34.0563 | -5.6829
3.1498 | -34.0566 | -5.6948
3.1559 | -34.0569 | -5.7062
3.162 | -34.0572 | -5.7181
3.1681 | -34.0574 | -5.73
3.1742 | -34.0578 | -5.7414
3.1803 | -34.0581 | -5.753
3.1864 | -34.0584 | -5.7641
3.1925 | -34.0587 | -5.7756
3.1986 | -34.0589 | -5.7874
3.2047 | -34.0592 | -5.7993
3.2108 | -34.0595 | -5.8103
3.2169 | -34.0598 | -5.8219
3.223 | -34.06 | -5.8335
3.2292 | -34.0602 | -5.8452
3.2353 | -34.0604 | -5.8559
3.2414 | -34.0606 | -5.8672
3.2475 | -34.0608 | -5.8786
3.2536 | -34.0611 | -5.8899
3.2597 | -34.0614 | -5.9007
3.2658 | -34.0616 | -5.9116
3.2719 | -34.0618 | -5.9225
3.278 | -34.062 | -5.934
3.2841 | -34.0623 | -5.945
3.2902 | -34.0625 | -5.9563
3.2963 | -34.0627 | -5.9677
3.3024 | -34.0628 | -5.9787
3.3085 | -34.0631 | -5.9901
3.3146 | -34.0633 | -6.0013
3.3207 | -34.0635 | -6.0127
3.3268 | -34.0637 | -6.024
3.3329 | -34.0639 | -6.0346
3.339 | -34.064 | -6.0454
3.3451 | -34.0642 | -6.0565
3.3512 | -34.0644 | -6.0671
3.3573 | -34.0646 | -6.0782
3.3634 | -34.0648 | -6.0893
3.3696 | -34.065 | -6.1005
3.3757 | -34.0651 | -6.1112
3.3818 | -34.0653 | -6.122
3.3879 | -34.0655 | -6.133
3.394 | -34.0656 | -6.1444
3.4001 | -34.0658 | -6.1551
3.4062 | -34.0659 | -6.166
3.4123 | -34.0661 | -6.1765
3.4184 | -34.0663 | -6.1877
3.4245 | -34.0665 | -6.1984
3.4306 | -34.0665 | -6.2091
3.4367 | -34.0667 | -6.2195
3.4428 | -34.0669 | -6.2299
3.4489 | -34.067 | -6.2408
3.455 | -34.0672 | -6.2516
3.4611 | -34.0674 | -6.2627
3.4672 | -34.0675 | -6.2738
3.4733 | -34.0677 | -6.2848
3.4794 | -34.0678 | -6.2956
3.4855 | -34.068 | -6.306
3.4916 | -34.0682 | -6.3163
3.4977 | -34.0683 | -6.3275
3.5038 | -34.0685 | -6.3383
3.5099 | -34.0686 | -6.3486
3.5161 | -34.0687 | -6.3589
3.5222 | -34.0689 | -6.3693
3.5283 | -34.0689 | -6.38
3.5344 | -34.0691 | -6.3908
3.5405 | -34.0692 | -6.4013
3.5466 | -34.0694 | -6.4122
3.5527 | -34.0695 | -6.423
3.5588 | -34.0696 | -6.4336
3.5649 | -34.0697 | -6.4443
3.571 | -34.0698 | -6.4549
3.5771 | -34.0699 | -6.4659
3.5832 | -34.0701 | -6.4758
3.5893 | -34.0703 | -6.4864
3.5954 | -34.0705 | -6.4972
3.6015 | -34.0706 | -6.5081
3.6076 | -34.0709 | -6.5185
3.6137 | -34.071 | -6.5286
3.6198 | -34.0713 | -6.5389
3.6259 | -34.0714 | -6.5498
3.632 | -34.0716 | -6.5604
3.6381 | -34.0717 | -6.5711
3.6442 | -34.0719 | -6.5817
3.6503 | -34.0721 | -6.5918
3.6565 | -34.0723 | -6.6026
3.6626 | -34.0725 | -6.6131
3.6687 | -34.0727 | -6.6238
3.6748 | -34.073 | -6.6341
3.6809 | -34.0732 | -6.6448
3.687 | -34.0734 | -6.6557
3.6931 | -34.0736 | -6.6659
3.6992 | -34.0738 | -6.6766
3.7053 | -34.074 | -6.6871
3.7114 | -34.0742 | -6.6978
3.7175 | -34.0744 | -6.7084
3.7236 | -34.0747 | -6.7185
3.7297 | -34.0749 | -6.7292
3.7358 | -34.0751 | -6.7401
3.7419 | -34.0753 | -6.7509
3.748 | -34.0755 | -6.7614
3.7541 | -34.0757 | -6.772
3.7602 | -34.076 | -6.7826
3.7663 | -34.0762 | -6.793
3.7724 | -34.0764 | -6.8036
3.7785 | -34.0767 | -6.8141
3.7846 | -34.0769 | -6.8245
3.7907 | -34.0772 | -6.8354
3.7969 | -34.0775 | -6.8461
3.803 | -34.0777 | -6.8563
3.8091 | -34.0779 | -6.8668
3.8152 | -34.0782 | -6.8773
3.8213 | -34.0784 | -6.8882
3.8274 | -34.0787 | -6.8985
3.8335 | -34.0791 | -6.9092
3.8396 | -34.0793 | -6.9198
3.8457 | -34.0796 | -6.9306
3.8518 | -34.08 | -6.941
3.8579 | -34.0803 | -6.9516
3.864 | -34.0806 | -6.9623
3.8701 | -34.0809 | -6.9733
3.8762 | -34.0811 | -6.9837
3.8823 | -34.0814 | -6.9945
3.8884 | -34.0817 | -7.0051
3.8945 | -34.082 | -7.016
3.9006 | -34.0823 | -7.027
3.9067 | -34.0826 | -7.0381
3.9128 | -34.0829 | -7.0488
3.9189 | -34.0832 | -7.0593
3.925 | -34.0834 | -7.0701
3.9311 | -34.0837 | -7.0803
3.9372 | -34.084 | -7.0911
3.9434 | -34.0843 | -7.1017
3.9495 | -34.0846 | -7.1125
3.9556 | -34.0849 | -7.123
3.9617 | -34.0852 | -7.1332
3.9678 | -34.0856 | -7.1436
3.9739 | -34.0858 | -7.1542
3.98 | -34.0861 | -7.1649
3.9861 | -34.0863 | -7.1755
3.9922 | -34.0866 | -7.1862
3.9983 | -34.0869 | -7.1964
4.0044 | -34.0871 | -7.2066
4.0105 | -34.0874 | -7.2173
4.0166 | -34.0877 | -7.228
4.0227 | -34.088 | -7.2385
4.0288 | -34.0883 | -7.2496
4.0349 | -34.0885 | -7.26
4.041 | -34.0888 | -7.2705
4.0471 | -34.089 | -7.2812
4.0532 | -34.0893 | -7.2915
4.0593 | -34.0896 | -7.3022
4.0654 | -34.0898 | -7.3126
4.0715 | -34.0901 | -7.3237
4.0776 | -34.0904 | -7.3342
4.0838 | -34.0907 | -7.3453
4.0899 | -34.091 | -7.3556
4.096 | -34.0913 | -7.3662
4.1021 | -34.0915 | -7.3765
4.1082 | -34.0918 | -7.3869
4.1143 | -34.092 | -7.3971
4.1204 | -34.0923 | -7.4074
4.1265 | -34.0926 | -7.4177
4.1326 | -34.0928 | -7.4279
4.1387 | -34.0931 | -7.4384
4.1448 | -34.0934 | -7.449
4.1509 | -34.0936 | -7.4596
4.157 | -34.0938 | -7.4705
4.1631 | -34.0941 | -7.4812
4.1692 | -34.0944 | -7.4916
4.1753 | -34.0946 | -7.5021
4.1814 | -34.0948 | -7.5123
4.1875 | -34.0951 | -7.5231
4.1936 | -34.0953 | -7.5337
4.1997 | -34.0956 | -7.5438
4.2058 | -34.0958 | -7.5546
4.2119 | -34.096 | -7.5649
4.218 | -34.0962 | -7.5751
4.2241 | -34.0965 | -7.5854
4.2303 | -34.0967 | -7.5957
4.2364 | -34.097 | -7.6061
4.2425 | -34.0971 | -7.6166
4.2486 | -34.0973 | -7.6272
4.2547 | -34.0975 | -7.6374
4.2608 | -34.0977 | -7.648
4.2669 | -34.0979 | -7.6581
4.273 | -34.0981 | -7.6687
4.2791 | -34.0983 | -7.6791
4.2852 | -34.0986 | -7.6897
4.2913 | -34.0987 | -7.6998
4.2974 | -34.0989 | -7.7105
4.3035 | -34.0992 | -7.7212
4.3096 | -34.0993 | -7.7318
4.3157 | -34.0996 | -7.7421
4.3218 | -34.0998 | -7.7524
4.3279 | -34.1 | -7.7628
4.334 | -34.1001 | -7.7731
4.3401 | -34.1003 | -7.7831
4.3462 | -34.1006 | -7.7936
4.3523 | -34.1009 | -7.8035
4.3584 | -34.101 | -7.8135
4.3645 | -34.1012 | -7.8238
4.3707 | -34.1014 | -7.8338
4.3768 | -34.1016 | -7.8441
4.3829 | -34.1018 | -7.8548
4.389 | -34.102 | -7.8651
4.3951 | -34.1022 | -7.8752
4.4012 | -34.1024 | -7.8858
4.4073 | -34.1026 | -7.8959
4.4134 | -34.1028 | -7.9062
4.4195 | -34.103 | -7.9165
4.4256 | -34.1033 | -7.9266
4.4317 | -34.1035 | -7.9368
4.4378 | -34.1037 | -7.9468
4.4439 | -34.1039 | -7.9567
4.45 | -34.104 | -7.9669
4.4561 | -34.1041 | -7.9772
4.4622 | -34.1043 | -7.9875
4.4683 | -34.1045 | -7.9977
4.4744 | -34.1047 | -8.0083
4.4805 | -34.1049 | -8.0181
4.4866 | -34.1052 | -8.0283
4.4927 | -34.1054 | -8.0388
4.4988 | -34.1056 | -8.0489
4.5049 | -34.1058 | -8.0588
4.511 | -34.106 | -8.0687
4.5172 | -34.1062 | -8.079
4.5233 | -34.1064 | -8.0896
4.5294 | -34.1066 | -8.0996
4.5355 | -34.1069 | -8.1098
4.5416 | -34.1071 | -8.12
4.5477 | -34.1073 | -8.1302
4.5538 | -34.1075 | -8.14
4.5599 | -34.1077 | -8.1502
4.566 | -34.1079 | -8.1604
4.5721 | -34.1081 | -8.1706
4.5782 | -34.1084 | -8.181
4.5843 | -34.1086 | -8.1913
4.5904 | -34.1089 | -8.2016
4.5965 | -34.1091 | -8.2116
4.6026 | -34.1093 | -8.2222
4.6087 | -34.1095 | -8.2323
4.6148 | -34.1097 | -8.2429
4.6209 | -34.1099 | -8.253
4.627 | -34.1102 | -8.2634
4.6331 | -34.1103 | -8.2735
4.6392 | -34.1106 | -8.2838
4.6453 | -34.1108 | -8.294
4.6514 | -34.111 | -8.3042
4.6576 | -34.1112 | -8.3148
4.6637 | -34.1114 | -8.3247
4.6698 | -34.1116 | -8.3348
4.6759 | -34.1118 | -8.3454
4.682 | -34.112 | -8.3556
4.6881 | -34.1122 | -8.3658
4.6942 | -34.1124 | -8.3759
4.7003 | -34.1126 | -8.3862
4.7064 | -34.1129 | -8.3963
4.7125 | -34.1131 | -8.4069
4.7186 | -34.1133 | -8.4169
4.7247 | -34.1136 | -8.4272
4.7308 | -34.1138 | -8.4374
4.7369 | -34.114 | -8.4478
4.743 | -34.1143 | -8.4584
4.7491 | -34.1145 | -8.469
4.7552 | -34.1147 | -8.4794
4.7613 | -34.1149 | -8.4898
4.7674 | -34.1151 | -8.4999
4.7735 | -34.1153 | -8.5106
4.7796 | -34.1156 | -8.5212
4.7857 | -34.1158 | -8.5316
4.7918 | -34.1161 | -8.5417
4.7979 | -34.1164 | -8.5519
4.8041 | -34.1166 | -8.5623
4.8102 | -34.1168 | -8.5722
4.8163 | -34.1171 | -8.5825
4.8224 | -34.1173 | -8.5928
4.8285 | -34.1176 | -8.6028
4.8346 | -34.1178 | -8.6129
4.8407 | -34.118 | -8.6233
4.8468 | -34.1183 | -8.6339
4.8529 | -34.1185 | -8.6446
4.859 | -34.1188 | -8.6552
4.8651 | -34.119 | -8.6653
4.8712 | -34.1193 | -8.6755
4.8773 | -34.1195 | -8.6859
4.8834 | -34.1197 | -8.6962
4.8895 | -34.1199 | -8.7063
4.8956 | -34.1201 | -8.7167
4.9017 | -34.1203 | -8.7269
4.9078 | -34.1206 | -8.7375
4.9139 | -34.1208 | -8.7475
4.92 | -34.1212 | -8.7575
4.9261 | -34.1214 | -8.7675
4.9322 | -34.1217 | -8.7779
4.9383 | -34.122 | -8.7881
4.9445 | -34.1222 | -8.7985
4.9506 | -34.1224 | -8.8086
4.9567 | -34.1227 | -8.8188
4.9628 | -34.1229 | -8.8291
4.9689 | -34.1232 | -8.8398
4.975 | -34.1234 | -8.8504
4.9811 | -34.1237 | -8.8604
4.9872 | -34.124 | -8.8706
4.9933 | -34.1243 | -8.8813
4.9994 | -34.1246 | -8.892
5.0055 | -34.1249 | -8.9024
5.0116 | -34.1251 | -8.9131
5.0177 | -34.1254 | -8.9238
5.0238 | -34.1257 | -8.9345
5.0299 | -34.1261 | -8.9451
5.036 | -34.1264 | -8.9555
5.0421 | -34.1266 | -8.9657
5.0482 | -34.1269 | -8.9761
5.0543 | -34.1271 | -8.9863
5.0604 | -34.1275 | -8.9967
5.0665 | -34.1278 | -9.0067
5.0726 | -34.128 | -9.0175
5.0787 | -34.1283 | -9.028
5.0848 | -34.1286 | -9.0386
5.091 | -34.1289 | -9.0488
5.0971 | -34.1292 | -9.0588
5.1032 | -34.1295 | -9.0695
5.1093 | -34.1298 | -9.0799
5.1154 | -34.13 | -9.0901
5.1215 | -34.1303 | -9.1003
5.1276 | -34.1306 | -9.1105
5.1337 | -34.1309 | -9.1206
5.1398 | -34.1312 | -9.1307
5.1459 | -34.1315 | -9.1414
5.152 | -34.1318 | -9.1519
5.1581 | -34.132 | -9.1625
5.1642 | -34.1324 | -9.1729
5.1703 | -34.1326 | -9.1832
5.1764 | -34.133 | -9.1935
5.1825 | -34.1332 | -9.2036
5.1886 | -34.1335 | -9.2136
5.1947 | -34.1338 | -9.224
5.2008 | -34.1341 | -9.2345
5.2069 | -34.1344 | -9.245
5.213 | -34.1347 | -9.2552
5.2191 | -34.135 | -9.2656
5.2252 | -34.1353 | -9.276
5.2314 | -34.1356 | -9.2866
5.2375 | -34.1359 | -9.2967
5.2436 | -34.1361 | -9.3071
5.2497 | -34.1364 | -9.3172
5.2558 | -34.1368 | -9.3279
5.2619 | -34.137 | -9.3385
5.268 | -34.1374 | -9.3491
5.2741 | -34.1377 | -9.3595
5.2802 | -34.138 | -9.3702
5.2863 | -34.1384 | -9.3805
5.2924 | -34.1387 | -9.3911
5.2985 | -34.139 | -9.4016
5.3046 | -34.1393 | -9.412
5.3107 | -34.1396 | -9.4226
5.3168 | -34.1399 | -9.433
5.3229 | -34.1402 | -9.4434
5.329 | -34.1405 | -9.4539
5.3351 | -34.1408 | -9.464
5.3412 | -34.1411 | -9.4743
5.3473 | -34.1414 | -9.4847
5.3534 | -34.1417 | -9.4952
5.3595 | -34.142 | -9.5056
5.3656 | -34.1424 | -9.516
5.3717 | -34.1427 | -9.5267
5.3779 | -34.143 | -9.5374
5.384 | -34.1433 | -9.5478
5.3901 | -34.1436 | -9.5585
5.3962 | -34.1439 | -9.5687
5.4023 | -34.1442 | -9.5793
5.4084 | -34.1446 | -9.5896
5.4145 | -34.1449 | -9.5998
5.4206 | -34.1452 | -9.6102
5.4267 | -34.1455 | -9.621
5.4328 | -34.1458 | -9.6314
5.4389 | -34.1461 | -9.6415
5.445 | -34.1464 | -9.6522
5.4511 | -34.1467 | -9.6624
5.4572 | -34.1471 | -9.673
5.4633 | -34.1474 | -9.6834
5.4694 | -34.1476 | -9.6938
5.4755 | -34.1479 | -9.704
5.4816 | -34.1482 | -9.7144
5.4877 | -34.1486 | -9.7247
5.4938 | -34.1489 | -9.7348
5.4999 | -34.1492 | -9.7451
5.506 | -34.1496 | -9.7554
5.5121 | -34.1499 | -9.7657
5.5183 | -34.1502 | -9.7758
5.5244 | -34.1504 | -9.7864
5.5305 | -34.1508 | -9.7967
5.5366 | -34.1511 | -9.8068
5.5427 | -34.1514 | -9.8173
5.5488 | -34.1517 | -9.8277
5.5549 | -34.1519 | -9.838
5.561 | -34.1523 | -9.8486
5.5671 | -34.1526 | -9.8588
5.5732 | -34.1529 | -9.8694
5.5793 | -34.1532 | -9.88
5.5854 | -34.1536 | -9.8903
5.5915 | -34.1539 | -9.901
5.5976 | -34.1542 | -9.9114
5.6037 | -34.1545 | -9.9218
5.6098 | -34.1547 | -9.9321
5.6159 | -34.1551 | -9.9423
5.622 | -34.1554 | -9.9527
5.6281 | -34.1557 | -9.9631
5.6342 | -34.156 | -9.9738
5.6403 | -34.1563 | -9.9842
5.6464 | -34.1566 | -9.9947
5.6525 | -34.1569 | -10.0052
5.6586 | -34.1572 | -10.0152
5.6648 | -34.1576 | -10.0257
5.6709 | -34.1579 | -10.0362
5.677 | -34.1582 | -10.0466
5.6831 | -34.1585 | -10.057
5.6892 | -34.1588 | -10.0676
5.6953 | -34.1591 | -10.0779
5.7014 | -34.1594 | -10.0883
5.7075 | -34.1597 | -10.0987
5.7136 | -34.16 | -10.109
5.7197 | -34.1603 | -10.1193
5.7258 | -34.1606 | -10.1295
5.7319 | -34.1609 | -10.1398
5.738 | -34.1612 | -10.1502
5.7441 | -34.1615 | -10.1605
5.7502 | -34.1619 | -10.1706
5.7563 | -34.1622 | -10.1809
5.7624 | -34.1625 | -10.1915
5.7685 | -34.1628 | -10.2018
5.7746 | -34.1631 | -10.2125
5.7807 | -34.1633 | -10.2229
5.7868 | -34.1636 | -10.2334
5.7929 | -34.1639 | -10.2437
5.799 | -34.1642 | -10.2538
5.8052 | -34.1645 | -10.264
5.8113 | -34.1648 | -10.2742
5.8174 | -34.1651 | -10.2844
5.8235 | -34.1654 | -10.2947
5.8296 | -34.1657 | -10.3052
5.8357 | -34.166 | -10.3154
5.8418 | -34.1663 | -10.3258
5.8479 | -34.1665 | -10.3362
5.854 | -34.1668 | -10.3467
5.8601 | -34.1672 | -10.3569
5.8662 | -34.1675 | -10.3674
5.8723 | -34.1677 | -10.3775
5.8784 | -34.168 | -10.3877
5.8845 | -34.1683 | -10.3979
5.8906 | -34.1686 | -10.4083
5.8967 | -34.1689 | -10.4184
5.9028 | -34.1693 | -10.4286
5.9089 | -34.1696 | -10.4387
5.915 | -34.1699 | -10.4488
5.9211 | -34.1702 | -10.459
5.9272 | -34.1705 | -10.4691
5.9333 | -34.1708 | -10.4792
5.9394 | -34.1711 | -10.4895
5.9455 | -34.1714 | -10.4995
5.9517 | -34.1718 | -10.5096
5.9578 | -34.1721 | -10.52
5.9639 | -34.1724 | -10.5299
5.97 | -34.1728 | -10.5403
5.9761 | -34.1731 | -10.5504
5.9822 | -34.1734 | -10.5608
5.9883 | -34.1737 | -10.5711
5.9944 | -34.174 | -10.5815
6.0005 | -34.1744 | -10.5919
6.0066 | -34.1747 | -10.6019
6.0127 | -34.175 | -10.6119
6.0188 | -34.1753 | -10.6224
6.0249 | -34.1756 | -10.6327
6.031 | -34.1759 | -10.6427
6.0371 | -34.1762 | -10.6529
6.0432 | -34.1765 | -10.6631
6.0493 | -34.1768 | -10.6732
6.0554 | -34.1772 | -10.6832
6.0615 | -34.1775 | -10.6935
6.0676 | -34.1778 | -10.7039
6.0737 | -34.1781 | -10.7142
6.0798 | -34.1784 | -10.7244
6.0859 | -34.1788 | -10.7346
6.0921 | -34.1791 | -10.7446
6.0982 | -34.1794 | -10.7548
6.1043 | -34.1797 | -10.7649
6.1104 | -34.1801 | -10.7753
6.1165 | -34.1804 | -10.7851
6.1226 | -34.1807 | -10.7952
6.1287 | -34.181 | -10.8054
6.1348 | -34.1813 | -10.8155
6.1409 | -34.1817 | -10.8257
6.147 | -34.182 | -10.8358
6.1531 | -34.1823 | -10.846
6.1592 | -34.1826 | -10.856
6.1653 | -34.1829 | -10.8661
6.1714 | -34.1832 | -10.8763
6.1775 | -34.1835 | -10.8866
6.1836 | -34.1838 | -10.8968
6.1897 | -34.1842 | -10.907
6.1958 | -34.1845 | -10.9174
6.2019 | -34.1848 | -10.9277
6.208 | -34.1852 | -10.9378
6.2141 | -34.1855 | -10.948
6.2202 | -34.1858 | -10.9582
6.2263 | -34.1862 | -10.9684
6.2325 | -34.1865 | -10.9783
6.2386 | -34.1868 | -10.9882
6.2447 | -34.1871 | -10.9985
6.2508 | -34.1874 | -11.0087
6.2569 | -34.1878 | -11.0188
6.263 | -34.1881 | -11.029
6.2691 | -34.1884 | -11.039
6.2752 | -34.1887 | -11.0493
6.2813 | -34.189 | -11.0594
6.2874 | -34.1893 | -11.0694
6.2935 | -34.1896 | -11.0795
6.2996 | -34.1899 | -11.0896
6.3057 | -34.1902 | -11.0998
6.3118 | -34.1905 | -11.1101
6.3179 | -34.1908 | -11.1202
6.324 | -34.1912 | -11.1304
6.3301 | -34.1915 | -11.1405
6.3362 | -34.1918 | -11.1506
6.3423 | -34.1922 | -11.1609
6.3484 | -34.1925 | -11.1711
6.3545 | -34.1928 | -11.181
6.3606 | -34.1931 | -11.1909
6.3667 | -34.1934 | -11.2009
6.3728 | -34.1937 | -11.211
6.379 | -34.194 | -11.2212
6.3851 | -34.1944 | -11.2311
6.3912 | -34.1946 | -11.2412
6.3973 | -34.1949 | -11.2514
6.4034 | -34.1953 | -11.2613
6.4095 | -34.1956 | -11.2715
6.4156 | -34.1959 | -11.2816
6.4217 | -34.1962 | -11.2917
6.4278 | -34.1965 | -11.3018
6.4339 | -34.1969 | -11.3119
6.44 | -34.1972 | -11.3219
6.4461 | -34.1975 | -11.3322
6.4522 | -34.1978 | -11.3421
6.4583 | -34.1981 | -11.3523
6.4644 | -34.1984 | -11.3625
6.4705 | -34.1987 | -11.3723
6.4766 | -34.199 | -11.3825
6.4827 | -34.1993 | -11.3926
6.4888 | -34.1996 | -11.4025
6.4949 | -34.1999 | -11.4123
6.501 | -34.2002 | -11.4226
6.5071 | -34.2006 | -11.4325
6.5132 | -34.2009 | -11.4426
6.5194 | -34.2012 | -11.4528
6.5255 | -34.2015 | -11.4626
6.5316 | -34.2018 | -11.4729
6.5377 | -34.2021 | -11.483
6.5438 | -34.2024 | -11.4929
6.5499 | -34.2027 | -11.5029
6.556 | -34.2031 | -11.5129
6.5621 | -34.2034 | -11.5232
6.5682 | -34.2037 | -11.5334
6.5743 | -34.204 | -11.5435
6.5804 | -34.2043 | -11.5537
6.5865 | -34.2047 | -11.5638
6.5926 | -34.205 | -11.5738
6.5987 | -34.2053 | -11.5836
6.6048 | -34.2056 | -11.5937
6.6109 | -34.2059 | -11.6039
6.617 | -34.2062 | -11.6141
6.6231 | -34.2065 | -11.6242
6.6292 | -34.2069 | -11.6343
6.6353 | -34.2072 | -11.6446
6.6414 | -34.2075 | -11.6546
6.6475 | -34.2078 | -11.6647
6.6536 | -34.2081 | -11.6748
6.6597 | -34.2084 | -11.6846
6.6659 | -34.2087 | -11.6947
6.672 | -34.2091 | -11.7047
6.6781 | -34.2094 | -11.7148
6.6842 | -34.2097 | -11.7247
6.6903 | -34.2101 | -11.7347
6.6964 | -34.2104 | -11.7448
6.7025 | -34.2107 | -11.7546
6.7086 | -34.211 | -11.7647
6.7147 | -34.2113 | -11.7748
6.7208 | -34.2116 | -11.785
6.7269 | -34.2119 | -11.795
6.733 | -34.2122 | -11.8048
6.7391 | -34.2126 | -11.8149
6.7452 | -34.2129 | -11.8247
6.7513 | -34.2132 | -11.8346
6.7574 | -34.2135 | -11.8447
6.7635 | -34.2138 | -11.8545
6.7696 | -34.2142 | -11.8644
6.7757 | -34.2145 | -11.8744
6.7818 | -34.2148 | -11.8845
6.7879 | -34.2152 | -11.8945
6.794 | -34.2155 | -11.9043
6.8001 | -34.2158 | -11.9143
6.8063 | -34.2161 | -11.9243
6.8124 | -34.2164 | -11.9345
6.8185 | -34.2167 | -11.9447
6.8246 | -34.217 | -11.9547
6.8307 | -34.2173 | -11.9646
6.8368 | -34.2176 | -11.9748
6.8429 | -34.218 | -11.9849
6.849 | -34.2183 | -11.995
6.8551 | -34.2187 | -12.0049
6.8612 | -34.219 | -12.0149
6.8673 | -34.2193 | -12.0249
6.8734 | -34.2196 | -12.0351
6.8795 | -34.2199 | -12.045
6.8856 | -34.2202 | -12.0551
6.8917 | -34.2206 | -12.065
6.8978 | -34.2209 | -12.0748
6.9039 | -34.2212 | -12.0845
6.91 | -34.2215 | -12.0946
6.9161 | -34.2219 | -12.1044
6.9222 | -34.2222 | -12.1144
6.9283 | -34.2225 | -12.1242
6.9344 | -34.2229 | -12.1341
6.9405 | -34.2232 | -12.1441
6.9466 | -34.2235 | -12.1543
6.9528 | -34.2239 | -12.1645
6.9589 | -34.2242 | -12.1746
6.965 | -34.2246 | -12.1847
6.9711 | -34.2249 | -12.1948
6.9772 | -34.2253 | -12.2047
6.9833 | -34.2256 | -12.2147
6.9894 | -34.2259 | -12.2249
6.9955 | -34.2263 | -12.2347
7.0016 | -34.2266 | -12.2447
7.0077 | -34.2269 | -12.2546
7.0138 | -34.2273 | -12.2645
7.0199 | -34.2276 | -12.2744
7.026 | -34.2279 | -12.2843
7.0321 | -34.2282 | -12.2941
7.0382 | -34.2286 | -12.3038
7.0443 | -34.2289 | -12.3139
7.0504 | -34.2292 | -12.3239
7.0565 | -34.2296 | -12.3339
7.0626 | -34.2299 | -12.3436
7.0687 | -34.2302 | -12.3532
7.0748 | -34.2306 | -12.3629
7.0809 | -34.2309 | -12.3728
7.087 | -34.2313 | -12.3829
7.0932 | -34.2316 | -12.3929
7.0993 | -34.2319 | -12.4028
7.1054 | -34.2323 | -12.4126
7.1115 | -34.2326 | -12.4223
7.1176 | -34.2329 | -12.4321
7.1237 | -34.2332 | -12.442
7.1298 | -34.2336 | -12.4518
7.1359 | -34.2339 | -12.4618
7.142 | -34.2342 | -12.4714
7.1481 | -34.2345 | -12.4812
7.1542 | -34.2349 | -12.4911
7.1603 | -34.2352 | -12.501
7.1664 | -34.2356 | -12.5109
7.1725 | -34.2359 | -12.5208
7.1786 | -34.2363 | -12.5306
7.1847 | -34.2366 | -12.5406
7.1908 | -34.2369 | -12.5507
7.1969 | -34.2372 | -12.5607
7.203 | -34.2376 | -12.5704
7.2091 | -34.2379 | -12.5802
7.2152 | -34.2383 | -12.5901
7.2213 | -34.2386 | -12.5997
7.2274 | -34.2389 | -12.6096
7.2335 | -34.2393 | -12.6194
7.2397 | -34.2396 | -12.6292
7.2458 | -34.24 | -12.6388
7.2519 | -34.2403 | -12.6487
7.258 | -34.2407 | -12.6587
7.2641 | -34.241 | -12.6685
7.2702 | -34.2413 | -12.6786
7.2763 | -34.2417 | -12.6885
7.2824 | -34.242 | -12.6984
7.2885 | -34.2424 | -12.7083
7.2946 | -34.2427 | -12.7181
7.3007 | -34.243 | -12.7281
7.3068 | -34.2433 | -12.7381
7.3129 | -34.2437 | -12.748
7.319 | -34.2441 | -12.7579
7.3251 | -34.2444 | -12.7675
7.3312 | -34.2447 | -12.7774
7.3373 | -34.2451 | -12.787
7.3434 | -34.2454 | -12.7968
7.3495 | -34.2458 | -12.8064
7.3556 | -34.2462 | -12.8164
7.3617 | -34.2465 | -12.8264
7.3678 | -34.2468 | -12.8362
7.3739 | -34.2472 | -12.8462
7.3801 | -34.2475 | -12.8562
7.3862 | -34.2479 | -12.866
7.3923 | -34.2482 | -12.8761
7.3984 | -34.2485 | -12.8857
7.4045 | -34.2489 | -12.8954
7.4106 | -34.2493 | -12.9052
7.4167 | -34.2496 | -12.9152
7.4228 | -34.25 | -12.9251
7.4289 | -34.2503 | -12.9349
7.435 | -34.2507 | -12.9445
7.4411 | -34.251 | -12.9543
7.4472 | -34.2514 | -12.9642
7.4533 | -34.2517 | -12.9738
7.4594 | -34.2521 | -12.9838
7.4655 | -34.2524 | -12.9935
7.4716 | -34.2528 | -13.0035
7.4777 | -34.2531 | -13.0133
7.4838 | -34.2535 | -13.023
7.4899 | -34.2538 | -13.0326
7.496 | -34.2541 | -13.0423
7.5021 | -34.2545 | -13.052
7.5082 | -34.2548 | -13.0619
7.5143 | -34.2552 | -13.0718
7.5204 | -34.2555 | -13.0817
7.5266 | -34.2558 | -13.0913
7.5327 | -34.2562 | -13.101
7.5388 | -34.2565 | -13.1111
7.5449 | -34.2569 | -13.1211
7.551 | -34.2573 | -13.1307
7.5571 | -34.2576 | -13.1406
7.5632 | -34.258 | -13.1505
7.5693 | -34.2583 | -13.1602
7.5754 | -34.2587 | -13.1699
7.5815 | -34.259 | -13.1797
7.5876 | -34.2594 | -13.1897
7.5937 | -34.2597 | -13.1996
7.5998 | -34.2601 | -13.2093
7.6059 | -34.2605 | -13.2194
7.612 | -34.2608 | -13.2292
7.6181 | -34.2612 | -13.2389
7.6242 | -34.2615 | -13.2489
7.6303 | -34.2619 | -13.2589
7.6364 | -34.2622 | -13.2688
7.6425 | -34.2626 | -13.2787
7.6486 | -34.2629 | -13.2885
7.6547 | -34.2633 | -13.2982
7.6608 | -34.2636 | -13.308
7.667 | -34.264 | -13.3178
7.6731 | -34.2643 | -13.3277
7.6792 | -34.2647 | -13.3378
7.6853 | -34.265 | -13.3477
7.6914 | -34.2654 | -13.3576
7.6975 | -34.2657 | -13.3675
7.7036 | -34.2661 | -13.3773
7.7097 | -34.2665 | -13.3873
7.7158 | -34.2668 | -13.3972
7.7219 | -34.2672 | -13.4068
7.728 | -34.2675 | -13.4167
7.7341 | -34.2679 | -13.4266
7.7402 | -34.2683 | -13.4365
7.7463 | -34.2686 | -13.4464
7.7524 | -34.269 | -13.4565
7.7585 | -34.2693 | -13.4663
7.7646 | -34.2697 | -13.4762
7.7707 | -34.27 | -13.4861
7.7768 | -34.2703 | -13.4961
7.7829 | -34.2707 | -13.5062
7.789 | -34.271 | -13.5163
7.7951 | -34.2714 | -13.5262
7.8012 | -34.2717 | -13.536
7.8073 | -34.2721 | -13.5461
7.8135 | -34.2725 | -13.5559
7.8196 | -34.2729 | -13.5659
7.8257 | -34.2732 | -13.576
7.8318 | -34.2735 | -13.586
7.8379 | -34.2739 | -13.5957
7.844 | -34.2743 | -13.6055
7.8501 | -34.2746 | -13.6156
7.8562 | -34.275 | -13.6257
7.8623 | -34.2754 | -13.6355
7.8684 | -34.2757 | -13.6454
7.8745 | -34.276 | -13.6554
7.8806 | -34.2764 | -13.6654
7.8867 | -34.2768 | -13.6751
7.8928 | -34.2771 | -13.6847
7.8989 | -34.2775 | -13.6946
7.905 | -34.2778 | -13.7046
7.9111 | -34.2782 | -13.7145
7.9172 | -34.2785 | -13.7244
7.9233 | -34.2789 | -13.7344
7.9294 | -34.2792 | -13.7444
7.9355 | -34.2796 | -13.7541
7.9416 | -34.2799 | -13.7641
7.9477 | -34.2802 | -13.7738
7.9539 | -34.2806 | -13.7836
7.96 | -34.281 | -13.7935
7.9661 | -34.2813 | -13.8035
7.9722 | -34.2817 | -13.8132
7.9783 | -34.282 | -13.8233
7.9844 | -34.2824 | -13.8331
7.9905 | -34.2828 | -13.8432
7.9966 | -34.2831 | -13.8532
8.0027 | -34.2834 | -13.8631
8.0088 | -34.2838 | -13.873
8.0149 | -34.2841 | -13.8829
8.021 | -34.2845 | -13.8928
8.0271 | -34.2849 | -13.9029
8.0332 | -34.2852 | -13.9127
8.0393 | -34.2856 | -13.9227
8.0454 | -34.2859 | -13.9328
8.0515 | -34.2863 | -13.9427
8.0576 | -34.2867 | -13.9528
8.0637 | -34.287 | -13.9626
8.0698 | -34.2874 | -13.9724
8.0759 | -34.2877 | -13.9822
8.082 | -34.2881 | -13.992
8.0881 | -34.2884 | -14.0019
8.0942 | -34.2888 | -14.0119
8.1004 | -34.2892 | -14.0219
8.1065 | -34.2895 | -14.0316
8.1126 | -34.2898 | -14.0416
8.1187 | -34.2902 | -14.0516
8.1248 | -34.2905 | -14.0613
8.1309 | -34.2909 | -14.071
8.137 | -34.2913 | -14.0807
8.1431 | -34.2916 | -14.0908
8.1492 | -34.292 | -14.1007
8.1553 | -34.2923 | -14.1107
8.1614 | -34.2927 | -14.1208
8.1675 | -34.293 | -14.1307
8.1736 | -34.2934 | -14.1407
8.1797 | -34.2937 | -14.1507
8.1858 | -34.2941 | -14.1607
8.1919 | -34.2944 | -14.1706
8.198 | -34.2948 | -14.1803
8.2041 | -34.2952 | -14.1902
8.2102 | -34.2955 | -14.2002
8.2163 | -34.2959 | -14.2099
8.2224 | -34.2962 | -14.2198
8.2285 | -34.2966 | -14.2294
8.2346 | -34.2969 | -14.2393
8.2408 | -34.2973 | -14.2491
8.2469 | -34.2977 | -14.259
8.253 | -34.298 | -14.2688
8.2591 | -34.2984 | -14.2788
8.2652 | -34.2988 | -14.2887
8.2713 | -34.2991 | -14.2987
8.2774 | -34.2995 | -14.3086
8.2835 | -34.2999 | -14.3188
8.2896 | -34.3002 | -14.3285
8.2957 | -34.3006 | -14.3385
8.3018 | -34.3009 | -14.3484
8.3079 | -34.3013 | -14.3582
8.314 | -34.3017 | -14.3681
8.3201 | -34.3021 | -14.3779
8.3262 | -34.3024 | -14.3878
8.3323 | -34.3027 | -14.3977
8.3384 | -34.3031 | -14.4077
8.3445 | -34.3035 | -14.4174
8.3506 | -34.3038 | -14.4273
8.3567 | -34.3042 | -14.4372
8.3628 | -34.3045 | -14.4472
8.3689 | -34.3049 | -14.457
8.375 | -34.3052 | -14.4669
8.3812 | -34.3056 | -14.4766
8.3873 | -34.306 | -14.4865
8.3934 | -34.3063 | -14.4964
8.3995 | -34.3067 | -14.5062
8.4056 | -34.307 | -14.516
8.4117 | -34.3074 | -14.5259
8.4178 | -34.3077 | -14.5358
8.4239 | -34.3081 | -14.5455
8.43 | -34.3084 | -14.5554
8.4361 | -34.3088 | -14.5653
8.4422 | -34.3091 | -14.5751
8.4483 | -34.3095 | -14.5849
8.4544 | -34.3099 | -14.5948
8.4605 | -34.3102 | -14.6045
8.4666 | -34.3106 | -14.6145
8.4727 | -34.3109 | -14.6245
8.4788 | -34.3113 | -14.6343
8.4849 | -34.3116 | -14.6442
8.491 | -34.312 | -14.654
8.4971 | -34.3124 | -14.6637
8.5032 | -34.3127 | -14.6734
8.5093 | -34.3131 | -14.6831
8.5154 | -34.3135 | -14.693
8.5215 | -34.3139 | -14.7027
8.5277 | -34.3142 | -14.7125
8.5338 | -34.3146 | -14.7223
8.5399 | -34.3149 | -14.7321
8.546 | -34.3153 | -14.7418
8.5521 | -34.3157 | -14.7516
8.5582 | -34.316 | -14.7615
8.5643 | -34.3164 | -14.7711
8.5704 | -34.3167 | -14.7807
8.5765 | -34.3171 | -14.7906
8.5826 | -34.3175 | -14.8001
8.5887 | -34.3178 | -14.8098
8.5948 | -34.3182 | -14.8197
8.6009 | -34.3186 | -14.8293
8.607 | -34.319 | -14.8392
8.6131 | -34.3193 | -14.849
8.6192 | -34.3197 | -14.8587
8.6253 | -34.3201 | -14.8684
8.6314 | -34.3204 | -14.8781
8.6375 | -34.3208 | -14.888
8.6436 | -34.3212 | -14.8977
8.6497 | -34.3215 | -14.9075
8.6558 | -34.3219 | -14.9172
8.6619 | -34.3222 | -14.927
8.6681 | -34.3226 | -14.937
8.6742 | -34.323 | -14.9467
8.6803 | -34.3233 | -14.9563
8.6864 | -34.3237 | -14.9658
8.6925 | -34.3241 | -14.9754
8.6986 | -34.3244 | -14.9852
8.7047 | -34.3248 | -14.9951
8.7108 | -34.3252 | -15.0049
8.7169 | -34.3255 | -15.0147
8.723 | -34.3259 | -15.0243
8.7291 | -34.3263 | -15.0339
8.7352 | -34.3266 | -15.0438
8.7413 | -34.327 | -15.0536
8.7474 | -34.3273 | -15.0632
8.7535 | -34.3277 | -15.073
8.7596 | -34.3281 | -15.0827
8.7657 | -34.3285 | -15.0926
8.7718 | -34.3288 | -15.1025
8.7779 | -34.3292 | -15.1123
8.784 | -34.3296 | -15.122
8.7901 | -34.33 | -15.1318
8.7962 | -34.3303 | -15.1413
8.8023 | -34.3307 | -15.151
8.8084 | -34.3311 | -15.1607
8.8146 | -34.3314 | -15.1704
8.8207 | -34.3318 | -15.1801
8.8268 | -34.3322 | -15.1895
8.8329 | -34.3325 | -15.1992
8.839 | -34.3329 | -15.2089
8.8451 | -34.3333 | -15.2185
8.8512 | -34.3337 | -15.2281
8.8573 | -34.334 | -15.2379
8.8634 | -34.3344 | -15.2474
8.8695 | -34.3348 | -15.2573
8.8756 | -34.3351 | -15.267
8.8817 | -34.3355 | -15.2766
8.8878 | -34.3359 | -15.2865
8.8939 | -34.3363 | -15.2962
8.9 | -34.3366 | -15.3059
8.9061 | -34.337 | -15.3155
8.9122 | -34.3374 | -15.3252
8.9183 | -34.3378 | -15.3349
8.9244 | -34.3382 | -15.3448
8.9305 | -34.3386 | -15.3542
8.9366 | -34.339 | -15.3639
8.9427 | -34.3393 | -15.3736
8.9488 | -34.3397 | -15.3833
8.955 | -34.3401 | -15.3929
8.9611 | -34.3405 | -15.4028
8.9672 | -34.3409 | -15.4124
8.9733 | -34.3413 | -15.4219
8.9794 | -34.3417 | -15.4315
8.9855 | -34.3421 | -15.4413
8.9916 | -34.3424 | -15.4511
8.9977 | -34.3428 | -15.4608
9.0038 | -34.3432 | -15.4703
9.0099 | -34.3436 | -15.48
9.016 | -34.344 | -15.4897
9.0221 | -34.3444 | -15.4992
9.0282 | -34.3448 | -15.5087
9.0343 | -34.3451 | -15.5184
9.0404 | -34.3456 | -15.528
9.0465 | -34.3459 | -15.5378
9.0526 | -34.3463 | -15.5474
9.0587 | -34.3468 | -15.5571
9.0648 | -34.3472 | -15.5666
9.0709 | -34.3475 | -15.5763
9.077 | -34.3479 | -15.5859
9.0831 | -34.3483 | -15.5957
9.0892 | -34.3487 | -15.6054
9.0953 | -34.3491 | -15.6152
9.1015 | -34.3495 | -15.6248
9.1076 | -34.3499 | -15.6343
9.1137 | -34.3503 | -15.644
9.1198 | -34.3507 | -15.6535
9.1259 | -34.3511 | -15.6631
9.132 | -34.3515 | -15.6729
9.1381 | -34.3519 | -15.6824
9.1442 | -34.3523 | -15.6917
9.1503 | -34.3527 | -15.7015
9.1564 | -34.3531 | -15.7111
9.1625 | -34.3535 | -15.7208
9.1686 | -34.3539 | -15.7303
9.1747 | -34.3543 | -15.74
9.1808 | -34.3547 | -15.7496
9.1869 | -34.3551 | -15.7592
9.193 | -34.3555 | -15.7689
9.1991 | -34.3559 | -15.7786
9.2052 | -34.3563 | -15.7882
9.2113 | -34.3567 | -15.7978
9.2174 | -34.3572 | -15.8075
9.2235 | -34.3576 | -15.8171
9.2296 | -34.358 | -15.8267
9.2357 | -34.3584 | -15.8365
9.2419 | -34.3588 | -15.846
9.248 | -34.3592 | -15.8557
9.2541 | -34.3596 | -15.8654
9.2602 | -34.3601 | -15.8751
9.2663 | -34.3605 | -15.8844
9.2724 | -34.3609 | -15.894
9.2785 | -34.3613 | -15.9036
9.2846 | -34.3617 | -15.9131
9.2907 | -34.3621 | -15.9226
9.2968 | -34.3625 | -15.9324
9.3029 | -34.3629 | -15.9421
9.309 | -34.3633 | -15.9517
9.3151 | -34.3637 | -15.9611
9.3212 | -34.3642 | -15.9706
9.3273 | -34.3646 | -15.9802
9.3334 | -34.365 | -15.9897
9.3395 | -34.3654 | -15.9991
9.3456 | -34.3658 | -16.0086
9.3517 | -34.3662 | -16.0183
9.3578 | -34.3666 | -16.0279
9.3639 | -34.367 | -16.0373
9.37 | -34.3674 | -16.0469
9.3761 | -34.3678 | -16.0567
9.3822 | -34.3683 | -16.0663
9.3884 | -34.3687 | -16.0756
9.3945 | -34.3691 | -16.0853
9.4006 | -34.3695 | -16.0949
9.4067 | -34.3699 | -16.1045
9.4128 | -34.3704 | -16.1142
9.4189 | -34.3708 | -16.1239
9.425 | -34.3712 | -16.1335
9.4311 | -34.3716 | -16.143
9.4372 | -34.3721 | -16.1525
9.4433 | -34.3725 | -16.1621
9.4494 | -34.3729 | -16.1715
9.4555 | -34.3733 | -16.1812
9.4616 | -34.3738 | -16.1907
9.4677 | -34.3742 | -16.2003
9.4738 | -34.3746 | -16.2099
9.4799 | -34.375 | -16.2195
9.486 | -34.3754 | -16.2291
9.4921 | -34.3758 | -16.2386
9.4982 | -34.3763 | -16.248
9.5043 | -34.3767 | -16.2575
9.5104 | -34.3771 | -16.2669
9.5165 | -34.3775 | -16.2765
9.5226 | -34.3779 | -16.2862
9.5288 | -34.3784 | -16.2956
9.5349 | -34.3788 | -16.3052
9.541 | -34.3792 | -16.3146
9.5471 | -34.3796 | -16.3241
9.5532 | -34.38 | -16.3338
9.5593 | -34.3805 | -16.3434
9.5654 | -34.3809 | -16.3529
9.5715 | -34.3813 | -16.3624
9.5776 | -34.3817 | -16.3722
9.5837 | -34.3821 | -16.3816
9.5898 | -34.3825 | -16.3913
9.5959 | -34.3829 | -16.401
9.602 | -34.3833 | -16.4104
9.6081 | -34.3838 | -16.4198
9.6142 | -34.3842 | -16.4293
9.6203 | -34.3846 | -16.439
9.6264 | -34.385 | -16.4484
9.6325 | -34.3854 | -16.4581
9.6386 | -34.3858 | -16.4676
9.6447 | -34.3862 | -16.4771
9.6508 | -34.3867 | -16.4866
9.6569 | -34.3871 | -16.4961
9.663 | -34.3875 | -16.5058
9.6691 | -34.388 | -16.5153
9.6753 | -34.3884 | -16.5248
9.6814 | -34.3888 | -16.5341
9.6875 | -34.3892 | -16.5435
9.6936 | -34.3897 | -16.5528
9.6997 | -34.3901 | -16.5625
9.7058 | -34.3906 | -16.5719
9.7119 | -34.391 | -16.5815
9.718 | -34.3914 | -16.5912
9.7241 | -34.3918 | -16.6007
9.7302 | -34.3922 | -16.6102
9.7363 | -34.3927 | -16.6197
9.7424 | -34.3931 | -16.629
9.7485 | -34.3935 | -16.6386
9.7546 | -34.3939 | -16.648
9.7607 | -34.3944 | -16.6577
9.7668 | -34.3948 | -16.6672
9.7729 | -34.3952 | -16.6766
9.779 | -34.3957 | -16.6859
9.7851 | -34.3961 | -16.6956
9.7912 | -34.3966 | -16.705
9.7973 | -34.397 | -16.7146
9.8034 | -34.3974 | -16.7239
9.8095 | -34.3979 | -16.7336
9.8157 | -34.3983 | -16.7431
9.8218 | -34.3988 | -16.7528
9.8279 | -34.3992 | -16.7622
9.834 | -34.3996 | -16.7718
9.8401 | -34.4 | -16.7813
9.8462 | -34.4005 | -16.7905
9.8523 | -34.4009 | -16.8
9.8584 | -34.4013 | -16.8096
9.8645 | -34.4017 | -16.8193
9.8706 | -34.4021 | -16.8289
9.8767 | -34.4025 | -16.8383
9.8828 | -34.4029 | -16.8477
9.8889 | -34.4034 | -16.857
9.895 | -34.4038 | -16.8665
9.9011 | -34.4042 | -16.876
9.9072 | -34.4047 | -16.8855
9.9133 | -34.4051 | -16.8949
9.9194 | -34.4055 | -16.9043
9.9255 | -34.4059 | -16.9138
9.9316 | -34.4063 | -16.9233
9.9377 | -34.4068 | -16.9327
9.9438 | -34.4072 | -16.9423
9.9499 | -34.4076 | -16.9519
9.956 | -34.408 | -16.9614
9.9622 | -34.4085 | -16.971
9.9683 | -34.4089 | -16.9805
9.9744 | -34.4093 | -16.9901
9.9805 | -34.4097 | -16.9996
9.9866 | -34.4102 | -17.009
9.9927 | -34.4106 | -17.0184
9.9988 | -34.411 | -17.028
10.0049 | -34.4114 | -17.0374
10.011 | -34.4118 | -17.047
10.0171 | -34.4123 | -17.0563
10.0232 | -34.4127 | -17.0657
10.0293 | -34.4132 | -17.0752
10.0354 | -34.4136 | -17.0847
10.0415 | -34.414 | -17.0942
10.0476 | -34.4145 | -17.1038
10.0537 | -34.4149 | -17.1133
10.0598 | -34.4153 | -17.1228
10.0659 | -34.4157 | -17.1323
10.072 | -34.4161 | -17.1416
10.0781 | -34.4166 | -17.1509
10.0842 | -34.417 | -17.1604
10.0903 | -34.4174 | -17.1698
10.0964 | -34.4179 | -17.179
10.1026 | -34.4183 | -17.1885
10.1087 | -34.4187 | -17.198
10.1148 | -34.4191 | -17.2076
10.1209 | -34.4195 | -17.217
10.127 | -34.4199 | -17.2263
10.1331 | -34.4204 | -17.2356
10.1392 | -34.4208 | -17.2449
10.1453 | -34.4212 | -17.2543
10.1514 | -34.4216 | -17.2637
10.1575 | -34.4221 | -17.2733
10.1636 | -34.4225 | -17.2825
10.1697 | -34.4229 | -17.292
10.1758 | -34.4234 | -17.3013
10.1819 | -34.4238 | -17.3107
10.188 | -34.4242 | -17.3202
10.1941 | -34.4246 | -17.3296
10.2002 | -34.425 | -17.339
10.2063 | -34.4254 | -17.3484
10.2124 | -34.4258 | -17.3579
10.2185 | -34.4263 | -17.3674
10.2246 | -34.4267 | -17.3767
10.2307 | -34.4271 | -17.3862
10.2368 | -34.4276 | -17.3953
10.2429 | -34.428 | -17.4047
10.2491 | -34.4284 | -17.4142
10.2552 | -34.4288 | -17.4236
10.2613 | -34.4292 | -17.4331
10.2674 | -34.4297 | -17.4425
10.2735 | -34.4301 | -17.4521
10.2796 | -34.4305 | -17.4615
10.2857 | -34.431 | -17.471
10.2918 | -34.4314 | -17.4806
10.2979 | -34.4318 | -17.49
10.304 | -34.4323 | -17.4993
10.3101 | -34.4326 | -17.5087
10.3162 | -34.4331 | -17.5183
10.3223 | -34.4335 | -17.5276
10.3284 | -34.4339 | -17.537
10.3345 | -34.4343 | -17.5463
10.3406 | -34.4348 | -17.5557
10.3467 | -34.4352 | -17.565
10.3528 | -34.4356 | -17.5743
10.3589 | -34.436 | -17.5837
10.365 | -34.4364 | -17.5933
10.3711 | -34.4368 | -17.6027
10.3772 | -34.4373 | -17.6122
10.3833 | -34.4377 | -17.6216
10.3895 | -34.4382 | -17.631
10.3956 | -34.4386 | -17.6404
10.4017 | -34.439 | -17.6496
10.4078 | -34.4394 | -17.659
10.4139 | -34.4398 | -17.6684
10.42 | -34.4403 | -17.6779
10.4261 | -34.4407 | -17.6873
10.4322 | -34.4411 | -17.6966
10.4383 | -34.4415 | -17.706
10.4444 | -34.4419 | -17.7152
10.4505 | -34.4423 | -17.7247
10.4566 | -34.4428 | -17.734
10.4627 | -34.4432 | -17.7433
10.4688 | -34.4436 | -17.7525
10.4749 | -34.444 | -17.762
10.481 | -34.4445 | -17.7714
10.4871 | -34.4449 | -17.7806
10.4932 | -34.4453 | -17.7899
10.4993 | -34.4457 | -17.7993
10.5054 | -34.4462 | -17.8086
10.5115 | -34.4466 | -17.818
10.5176 | -34.447 | -17.8272
10.5237 | -34.4475 | -17.8367
10.5298 | -34.4479 | -17.846
10.536 | -34.4483 | -17.8554
10.5421 | -34.4487 | -17.8646
10.5482 | -34.4491 | -17.874
10.5543 | -34.4496 | -17.8833
10.5604 | -34.45 | -17.8925
10.5665 | -34.4504 | -17.9017
10.5726 | -34.4509 | -17.911
10.5787 | -34.4513 | -17.9203
10.5848 | -34.4517 | -17.9294
10.5909 | -34.4521 | -17.9388
10.597 | -34.4525 | -17.9482
10.6031 | -34.453 | -17.9576
10.6092 | -34.4534 | -17.9669
10.6153 | -34.4538 | -17.9763
10.6214 | -34.4542 | -17.9854
10.6275 | -34.4546 | -17.9947
10.6336 | -34.4551 | -18.0041
10.6397 | -34.4555 | -18.0135
10.6458 | -34.4559 | -18.0227
10.6519 | -34.4563 | -18.0322
10.658 | -34.4567 | -18.0415
10.6641 | -34.4572 | -18.0509
10.6702 | -34.4576 | -18.0602
10.6764 | -34.458 | -18.0696
10.6825 | -34.4584 | -18.0788
10.6886 | -34.4589 | -18.088
10.6947 | -34.4593 | -18.0973
10.7008 | -34.4597 | -18.1066
10.7069 | -34.4602 | -18.116
10.713 | -34.4606 | -18.1253
10.7191 | -34.461 | -18.1347
10.7252 | -34.4614 | -18.1438
10.7313 | -34.4619 | -18.153
10.7374 | -34.4623 | -18.1624
10.7435 | -34.4627 | -18.1717
10.7496 | -34.4631 | -18.1808
10.7557 | -34.4635 | -18.1902
10.7618 | -34.464 | -18.1992
10.7679 | -34.4644 | -18.2085
10.774 | -34.4648 | -18.2179
10.7801 | -34.4652 | -18.2272
10.7862 | -34.4656 | -18.2363
10.7923 | -34.466 | -18.2457
10.7984 | -34.4665 | -18.2548
10.8045 | -34.4669 | -18.264
10.8106 | -34.4673 | -18.2733
10.8168 | -34.4677 | -18.2824
10.8229 | -34.4681 | -18.2917
10.829 | -34.4686 | -18.3009
10.8351 | -34.469 | -18.3099
10.8412 | -34.4694 | -18.3191
10.8473 | -34.4699 | -18.3282
10.8534 | -34.4703 | -18.3374
10.8595 | -34.4707 | -18.3466
10.8656 | -34.4711 | -18.3556
10.8717 | -34.4716 | -18.3648
10.8778 | -34.472 | -18.3741
10.8839 | -34.4724 | -18.3833
10.89 | -34.4728 | -18.3925
10.8961 | -34.4733 | -18.4018
10.9022 | -34.4737 | -18.4109
10.9083 | -34.4741 | -18.4201
10.9144 | -34.4745 | -18.4291
10.9205 | -34.475 | -18.4384
10.9266 | -34.4754 | -18.4476
10.9327 | -34.4758 | -18.4568
10.9388 | -34.4763 | -18.4658
10.9449 | -34.4767 | -18.475
10.951 | -34.4771 | -18.4842
10.9571 | -34.4775 | -18.4934
10.9633 | -34.478 | -18.5027
10.9694 | -34.4784 | -18.5121
10.9755 | -34.4788 | -18.5212
10.9816 | -34.4792 | -18.5305
10.9877 | -34.4797 | -18.5396
10.9938 | -34.4801 | -18.549
10.9999 | -34.4806 | -18.558
11.006 | -34.481 | -18.5669
11.0121 | -34.4814 | -18.5761
11.0182 | -34.4818 | -18.5852
11.0243 | -34.4823 | -18.5944
11.0304 | -34.4827 | -18.6036
11.0365 | -34.4831 | -18.6127
11.0426 | -34.4835 | -18.6217
11.0487 | -34.484 | -18.6309
11.0548 | -34.4844 | -18.64
11.0609 | -34.4849 | -18.6492
11.067 | -34.4853 | -18.6583
11.0731 | -34.4857 | -18.6674
11.0792 | -34.4862 | -18.6763
11.0853 | -34.4866 | -18.6855
11.0914 | -34.487 | -18.6948
11.0975 | -34.4875 | -18.7038
11.1037 | -34.4879 | -18.7129
11.1098 | -34.4883 | -18.7222
11.1159 | -34.4887 | -18.7311
11.122 | -34.4892 | -18.7402
11.1281 | -34.4896 | -18.7494
11.1342 | -34.49 | -18.7586
11.1403 | -34.4905 | -18.7678
11.1464 | -34.4909 | -18.777
11.1525 | -34.4913 | -18.7858
11.1586 | -34.4917 | -18.7951
11.1647 | -34.4922 | -18.8041
11.1708 | -34.4926 | -18.8132
11.1769 | -34.493 | -18.8224
11.183 | -34.4934 | -18.8313
11.1891 | -34.4939 | -18.8404
11.1952 | -34.4943 | -18.8495
11.2013 | -34.4948 | -18.8586
11.2074 | -34.4952 | -18.8677
11.2135 | -34.4956 | -18.877
11.2196 | -34.4961 | -18.8859
11.2257 | -34.4965 | -18.8951
11.2318 | -34.4969 | -18.9041
11.2379 | -34.4974 | -18.9133
11.244 | -34.4978 | -18.9225
11.2502 | -34.4983 | -18.9315
11.2563 | -34.4987 | -18.9405
11.2624 | -34.4992 | -18.9495
11.2685 | -34.4996 | -18.9586
11.2746 | -34.5 | -18.9675
11.2807 | -34.5005 | -18.9767
11.2868 | -34.5009 | -18.9859
11.2929 | -34.5014 | -18.9946
11.299 | -34.5018 | -19.0038
11.3051 | -34.5023 | -19.013
11.3112 | -34.5027 | -19.0222
11.3173 | -34.5031 | -19.0311
11.3234 | -34.5036 | -19.0399
11.3295 | -34.504 | -19.0491
11.3356 | -34.5045 | -19.0583
11.3417 | -34.5049 | -19.0672
11.3478 | -34.5053 | -19.0765
11.3539 | -34.5057 | -19.0856
11.36 | -34.5062 | -19.0944
11.3661 | -34.5066 | -19.1033
11.3722 | -34.5071 | -19.1122
11.3783 | -34.5075 | -19.1213
11.3844 | -34.5079 | -19.1304
11.3906 | -34.5084 | -19.1395
11.3967 | -34.5088 | -19.1483
11.4028 | -34.5093 | -19.1575
11.4089 | -34.5097 | -19.1664
11.415 | -34.5101 | -19.1754
11.4211 | -34.5106 | -19.1845
11.4272 | -34.511 | -19.1936
11.4333 | -34.5114 | -19.2025
11.4394 | -34.5118 | -19.2116
11.4455 | -34.5123 | -19.2205
11.4516 | -34.5127 | -19.2295
11.4577 | -34.5131 | -19.2386
11.4638 | -34.5136 | -19.2474
11.4699 | -34.514 | -19.2565
11.476 | -34.5145 | -19.2656
11.4821 | -34.5149 | -19.2745
11.4882 | -34.5154 | -19.2836
11.4943 | -34.5158 | -19.2927
11.5004 | -34.5162 | -19.3015
11.5065 | -34.5167 | -19.3104
11.5126 | -34.5171 | -19.3193
11.5187 | -34.5176 | -19.3283
11.5248 | -34.518 | -19.3375
11.5309 | -34.5185 | -19.3465
11.5371 | -34.5189 | -19.3553
11.5432 | -34.5193 | -19.3643
11.5493 | -34.5198 | -19.3735
11.5554 | -34.5202 | -19.3825
11.5615 | -34.5206 | -19.3917
11.5676 | -34.5211 | -19.4007
11.5737 | -34.5215 | -19.4095
11.5798 | -34.522 | -19.4183
11.5859 | -34.5224 | -19.4273
11.592 | -34.5228 | -19.4362
11.5981 | -34.5233 | -19.4451
11.6042 | -34.5237 | -19.4541
11.6103 | -34.5242 | -19.4629
11.6164 | -34.5246 | -19.4718
11.6225 | -34.5251 | -19.4808
11.6286 | -34.5255 | -19.4897
11.6347 | -34.5259 | -19.4988
11.6408 | -34.5264 | -19.5076
11.6469 | -34.5268 | -19.5166
11.653 | -34.5272 | -19.5255
11.6591 | -34.5277 | -19.5344
11.6652 | -34.5281 | -19.5432
11.6713 | -34.5286 | -19.552
11.6775 | -34.529 | -19.5607
11.6836 | -34.5295 | -19.5697
11.6897 | -34.5299 | -19.5787
11.6958 | -34.5304 | -19.5877
11.7019 | -34.5308 | -19.5967
11.708 | -34.5312 | -19.6056
11.7141 | -34.5317 | -19.6143
11.7202 | -34.5321 | -19.6233
11.7263 | -34.5326 | -19.6323
11.7324 | -34.533 | -19.6414
11.7385 | -34.5334 | -19.6503
11.7446 | -34.5339 | -19.6591
11.7507 | -34.5343 | -19.6677
11.7568 | -34.5348 | -19.6767
11.7629 | -34.5352 | -19.6857
11.769 | -34.5356 | -19.6945
11.7751 | -34.5361 | -19.7034
11.7812 | -34.5366 | -19.712
11.7873 | -34.537 | -19.7209
11.7934 | -34.5375 | -19.7298
11.7995 | -34.5379 | -19.7386
11.8056 | -34.5383 | -19.7474
11.8117 | -34.5388 | -19.7563
11.8178 | -34.5392 | -19.7649
11.824 | -34.5397 | -19.7737
11.8301 | -34.5401 | -19.7825
11.8362 | -34.5406 | -19.7913
11.8423 | -34.541 | -19.8002
11.8484 | -34.5414 | -19.8091
11.8545 | -34.5419 | -19.8178
11.8606 | -34.5423 | -19.8266
11.8667 | -34.5428 | -19.8356
11.8728 | -34.5432 | -19.8443
11.8789 | -34.5437 | -19.8532
11.885 | -34.5441 | -19.8621
11.8911 | -34.5446 | -19.8709
11.8972 | -34.5451 | -19.8796
11.9033 | -34.5455 | -19.8884
11.9094 | -34.5459 | -19.8973
11.9155 | -34.5464 | -19.9061
11.9216 | -34.5469 | -19.9148
11.9277 | -34.5473 | -19.9236
11.9338 | -34.5477 | -19.9325
11.9399 | -34.5482 | -19.9414
11.946 | -34.5486 | -19.9502
11.9521 | -34.5491 | -19.959
11.9582 | -34.5495 | -19.9676
11.9644 | -34.55 | -19.9763
11.9705 | -34.5504 | -19.9851
11.9766 | -34.5509 | -19.9939
11.9827 | -34.5513 | -20.0027
11.9888 | -34.5518 | -20.0114
11.9949 | -34.5523 | -20.02
12.001 | -34.5527 | -20.0288
12.0071 | -34.5532 | -20.0375
12.0132 | -34.5536 | -20.0462
12.0193 | -34.554 | -20.0551
12.0254 | -34.5545 | -20.0639
12.0315 | -34.5549 | -20.0724
12.0376 | -34.5554 | -20.0811
12.0437 | -34.5558 | -20.0898
12.0498 | -34.5563 | -20.0987
12.0559 | -34.5567 | -20.1073
12.062 | -34.5572 | -20.1157
12.0681 | -34.5577 | -20.1245
12.0742 | -34.5581 | -20.1331
12.0803 | -34.5586 | -20.142
12.0864 | -34.559 | -20.151
12.0925 | -34.5594 | -20.1598
12.0986 | -34.5599 | -20.1684
12.1047 | -34.5604 | -20.1771
12.1109 | -34.5608 | -20.1857
12.117 | -34.5613 | -20.1944
12.1231 | -34.5617 | -20.2031
12.1292 | -34.5622 | -20.2118
12.1353 | -34.5626 | -20.2204
12.1414 | -34.5631 | -20.2292
12.1475 | -34.5635 | -20.238
12.1536 | -34.564 | -20.2468
12.1597 | -34.5644 | -20.2555
12.1658 | -34.5649 | -20.2641
12.1719 | -34.5654 | -20.2725
12.178 | -34.5658 | -20.2812
12.1841 | -34.5663 | -20.29
12.1902 | -34.5667 | -20.2987
12.1963 | -34.5672 | -20.3075
12.2024 | -34.5677 | -20.3161
12.2085 | -34.5681 | -20.3249
12.2146 | -34.5686 | -20.3336
12.2207 | -34.5691 | -20.3423
12.2268 | -34.5695 | -20.351
12.2329 | -34.57 | -20.3596
12.239 | -34.5705 | -20.368
12.2451 | -34.5709 | -20.3767
12.2513 | -34.5714 | -20.3855
12.2574 | -34.5718 | -20.3942
12.2635 | -34.5723 | -20.4029
12.2696 | -34.5727 | -20.4115
12.2757 | -34.5732 | -20.4198
12.2818 | -34.5737 | -20.4284
12.2879 | -34.5741 | -20.4372
12.294 | -34.5746 | -20.4456
12.3001 | -34.575 | -20.4543
12.3062 | -34.5755 | -20.463
12.3123 | -34.576 | -20.4714
12.3184 | -34.5764 | -20.4799
12.3245 | -34.5769 | -20.4885
12.3306 | -34.5773 | -20.4971
12.3367 | -34.5778 | -20.5058
12.3428 | -34.5783 | -20.5145
12.3489 | -34.5787 | -20.5229
12.355 | -34.5792 | -20.5316
12.3611 | -34.5797 | -20.5404
12.3672 | -34.5801 | -20.549
12.3733 | -34.5806 | -20.5576
12.3794 | -34.581 | -20.566
12.3855 | -34.5815 | -20.5745
12.3916 | -34.5819 | -20.5832
12.3978 | -34.5824 | -20.5918
12.4039 | -34.5829 | -20.6005
12.41 | -34.5834 | -20.6091
12.4161 | -34.5838 | -20.6174
12.4222 | -34.5843 | -20.626
12.4283 | -34.5847 | -20.6345
12.4344 | -34.5852 | -20.6433
12.4405 | -34.5857 | -20.652
12.4466 | -34.5861 | -20.6605
12.4527 | -34.5866 | -20.669
12.4588 | -34.5871 | -20.6776
12.4649 | -34.5875 | -20.6862
12.471 | -34.588 | -20.6949
12.4771 | -34.5885 | -20.7035
12.4832 | -34.5889 | -20.712
12.4893 | -34.5894 | -20.7203
12.4954 | -34.5898 | -20.7288
12.5015 | -34.5903 | -20.7373
12.5076 | -34.5908 | -20.7459
12.5137 | -34.5912 | -20.7545
12.5198 | -34.5917 | -20.7629
12.5259 | -34.5922 | -20.7716
12.532 | -34.5926 | -20.7802
12.5382 | -34.5931 | -20.7888
12.5443 | -34.5935 | -20.7973
12.5504 | -34.594 | -20.806
12.5565 | -34.5944 | -20.8144
12.5626 | -34.5949 | -20.8232
12.5687 | -34.5954 | -20.8318
12.5748 | -34.5958 | -20.8404
12.5809 | -34.5963 | -20.849
12.587 | -34.5968 | -20.8576
12.5931 | -34.5972 | -20.8659
12.5992 | -34.5977 | -20.8746
12.6053 | -34.5982 | -20.8831
12.6114 | -34.5986 | -20.8915
12.6175 | -34.5991 | -20.9002
12.6236 | -34.5995 | -20.9088
12.6297 | -34.6 | -20.9172
12.6358 | -34.6005 | -20.9256
12.6419 | -34.601 | -20.9342
12.648 | -34.6014 | -20.9428
12.6541 | -34.6019 | -20.9514
12.6602 | -34.6024 | -20.9597
12.6663 | -34.6028 | -20.9683
12.6724 | -34.6033 | -20.9768
12.6785 | -34.6038 | -20.9853
12.6847 | -34.6042 | -20.9937
12.6908 | -34.6047 | -21.0022
12.6969 | -34.6052 | -21.0107
12.703 | -34.6056 | -21.0192
12.7091 | -34.6061 | -21.0278
12.7152 | -34.6066 | -21.0362
12.7213 | -34.607 | -21.0446
12.7274 | -34.6075 | -21.0532
12.7335 | -34.608 | -21.0615
12.7396 | -34.6084 | -21.0699
12.7457 | -34.6089 | -21.0785
12.7518 | -34.6093 | -21.0869
12.7579 | -34.6098 | -21.0955
12.764 | -34.6103 | -21.1041
12.7701 | -34.6108 | -21.1125
12.7762 | -34.6112 | -21.1209
12.7823 | -34.6117 | -21.1294
12.7884 | -34.6122 | -21.138
12.7945 | -34.6127 | -21.1466
12.8006 | -34.6132 | -21.155
12.8067 | -34.6136 | -21.1635
12.8128 | -34.6141 | -21.1718
12.8189 | -34.6146 | -21.1803
12.8251 | -34.615 | -21.1888
12.8312 | -34.6155 | -21.1974
12.8373 | -34.616 | -21.2058
12.8434 | -34.6164 | -21.2143
12.8495 | -34.6169 | -21.2227
12.8556 | -34.6174 | -21.231
12.8617 | -34.6178 | -21.2396
12.8678 | -34.6183 | -21.2482
12.8739 | -34.6188 | -21.2564
12.88 | -34.6192 | -21.2649
12.8861 | -34.6197 | -21.2734
12.8922 | -34.6201 | -21.2819
12.8983 | -34.6206 | -21.2904
12.9044 | -34.6211 | -21.2987
12.9105 | -34.6215 | -21.3071
12.9166 | -34.622 | -21.3156
12.9227 | -34.6225 | -21.3241
12.9288 | -34.6229 | -21.3325
12.9349 | -34.6234 | -21.341
12.941 | -34.6239 | -21.3494
12.9471 | -34.6244 | -21.3578
12.9532 | -34.6249 | -21.3662
12.9593 | -34.6253 | -21.3748
12.9654 | -34.6258 | -21.3835
12.9716 | -34.6263 | -21.3921
12.9777 | -34.6268 | -21.4005
12.9838 | -34.6272 | -21.4089
12.9899 | -34.6277 | -21.4174
12.996 | -34.6281 | -21.4258
13.0021 | -34.6286 | -21.4343
13.0082 | -34.6291 | -21.4427
13.0143 | -34.6296 | -21.4508
13.0204 | -34.63 | -21.4592
13.0265 | -34.6305 | -21.4676
13.0326 | -34.631 | -21.476
13.0387 | -34.6315 | -21.4846
13.0448 | -34.6319 | -21.493
13.0509 | -34.6324 | -21.5013
13.057 | -34.6329 | -21.5097
13.0631 | -34.6334 | -21.5181
13.0692 | -34.6339 | -21.5264
13.0753 | -34.6343 | -21.535
13.0814 | -34.6348 | -21.5435
13.0875 | -34.6353 | -21.5518
13.0936 | -34.6358 | -21.5603
13.0997 | -34.6362 | -21.5686
13.1058 | -34.6367 | -21.5769
13.112 | -34.6372 | -21.5853
13.1181 | -34.6377 | -21.5937
13.1242 | -34.6381 | -21.6022
13.1303 | -34.6386 | -21.6107
13.1364 | -34.6391 | -21.6191
13.1425 | -34.6395 | -21.6275
13.1486 | -34.64 | -21.636
13.1547 | -34.6405 | -21.6443
13.1608 | -34.6409 | -21.6527
13.1669 | -34.6414 | -21.661
13.173 | -34.6419 | -21.6696
13.1791 | -34.6423 | -21.6779
13.1852 | -34.6428 | -21.6864
13.1913 | -34.6433 | -21.6947
13.1974 | -34.6438 | -21.7032
13.2035 | -34.6442 | -21.7116
13.2096 | -34.6447 | -21.7199
13.2157 | -34.6452 | -21.7282
13.2218 | -34.6456 | -21.7366
13.2279 | -34.6461 | -21.7451
13.234 | -34.6466 | -21.7534
13.2401 | -34.6471 | -21.7619
13.2462 | -34.6475 | -21.7702
13.2524 | -34.648 | -21.7787
13.2585 | -34.6485 | -21.787
13.2646 | -34.649 | -21.7954
13.2707 | -34.6495 | -21.8038
13.2768 | -34.6499 | -21.8123
13.2829 | -34.6504 | -21.8208
13.289 | -34.6509 | -21.829
13.2951 | -34.6514 | -21.8374
13.3012 | -34.6518 | -21.8458
13.3073 | -34.6523 | -21.8543
13.3134 | -34.6528 | -21.8628
13.3195 | -34.6532 | -21.8711
13.3256 | -34.6537 | -21.8793
13.3317 | -34.6542 | -21.8876
13.3378 | -34.6546 | -21.896
13.3439 | -34.6551 | -21.9044
13.35 | -34.6556 | -21.913
13.3561 | -34.6561 | -21.9214
13.3622 | -34.6565 | -21.9296
13.3683 | -34.657 | -21.9379
13.3744 | -34.6575 | -21.9462
13.3805 | -34.658 | -21.9547
13.3866 | -34.6584 | -21.9629
13.3927 | -34.6589 | -21.9713
13.3989 | -34.6594 | -21.9797
13.405 | -34.6598 | -21.9881
13.4111 | -34.6603 | -21.9963
13.4172 | -34.6608 | -22.0048
13.4233 | -34.6612 | -22.0133
13.4294 | -34.6617 | -22.0218
13.4355 | -34.6622 | -22.03
13.4416 | -34.6627 | -22.0383
13.4477 | -34.6631 | -22.0467
13.4538 | -34.6636 | -22.0551
13.4599 | -34.6641 | -22.0636
13.466 | -34.6646 | -22.072
13.4721 | -34.665 | -22.0804
13.4782 | -34.6655 | -22.0887
13.4843 | -34.666 | -22.0972
13.4904 | -34.6665 | -22.1056
13.4965 | -34.6669 | -22.1141
13.5026 | -34.6674 | -22.1224
13.5087 | -34.6678 | -22.1308
13.5148 | -34.6683 | -22.1393
13.5209 | -34.6688 | -22.1479
13.527 | -34.6693 | -22.1562
13.5331 | -34.6698 | -22.1645
13.5393 | -34.6702 | -22.1728
13.5454 | -34.6707 | -22.1811
13.5515 | -34.6712 | -22.1896
13.5576 | -34.6716 | -22.1979
13.5637 | -34.6721 | -22.2064
13.5698 | -34.6726 | -22.2148
13.5759 | -34.6731 | -22.2229
13.582 | -34.6735 | -22.2312
13.5881 | -34.674 | -22.2396
13.5942 | -34.6745 | -22.2482
13.6003 | -34.675 | -22.2565
13.6064 | -34.6754 | -22.265
13.6125 | -34.6759 | -22.2734
13.6186 | -34.6764 | -22.2817
13.6247 | -34.6769 | -22.29
13.6308 | -34.6773 | -22.2985
13.6369 | -34.6778 | -22.3069
13.643 | -34.6783 | -22.3153
13.6491 | -34.6788 | -22.3236
13.6552 | -34.6792 | -22.3319
13.6613 | -34.6797 | -22.3403
13.6674 | -34.6802 | -22.3487
13.6735 | -34.6806 | -22.357
13.6796 | -34.6811 | -22.3653
13.6858 | -34.6816 | -22.3735
13.6919 | -34.6821 | -22.3819
13.698 | -34.6825 | -22.3903
13.7041 | -34.683 | -22.3988
13.7102 | -34.6835 | -22.4071
13.7163 | -34.684 | -22.4154
13.7224 | -34.6845 | -22.4239
13.7285 | -34.6849 | -22.4322
13.7346 | -34.6854 | -22.4406
13.7407 | -34.6859 | -22.449
13.7468 | -34.6864 | -22.4573
13.7529 | -34.6868 | -22.4657
13.759 | -34.6873 | -22.4742
13.7651 | -34.6878 | -22.4825
13.7712 | -34.6883 | -22.491
13.7773 | -34.6887 | -22.4995
13.7834 | -34.6892 | -22.5078
13.7895 | -34.6897 | -22.5161
13.7956 | -34.6901 | -22.5245
13.8017 | -34.6906 | -22.5328
13.8078 | -34.6911 | -22.5411
13.8139 | -34.6915 | -22.5494
13.82 | -34.692 | -22.5576
13.8262 | -34.6924 | -22.5661
13.8323 | -34.6929 | -22.5745
13.8384 | -34.6934 | -22.5829
13.8445 | -34.6939 | -22.5913
13.8506 | -34.6944 | -22.5997
13.8567 | -34.6949 | -22.6081
13.8628 | -34.6953 | -22.6164
13.8689 | -34.6958 | -22.6249
13.875 | -34.6963 | -22.6332
13.8811 | -34.6967 | -22.6415
13.8872 | -34.6972 | -22.6499
13.8933 | -34.6977 | -22.6582
13.8994 | -34.6982 | -22.6665
13.9055 | -34.6986 | -22.6751
13.9116 | -34.6991 | -22.6833
13.9177 | -34.6996 | -22.6917
13.9238 | -34.7001 | -22.7001
13.9299 | -34.7005 | -22.7084
13.936 | -34.701 | -22.7168
13.9421 | -34.7014 | -22.7251
13.9482 | -34.7019 | -22.7335
13.9543 | -34.7024 | -22.7419
13.9604 | -34.7028 | -22.7501
13.9665 | -34.7033 | -22.7584
13.9727 | -34.7038 | -22.7669
13.9788 | -34.7043 | -22.7751
13.9849 | -34.7048 | -22.7835
13.991 | -34.7053 | -22.7918
13.9971 | -34.7057 | -22.8
14.0032 | -34.7062 | -22.8084
14.0093 | -34.7067 | -22.8168
14.0154 | -34.7071 | -22.825
14.0215 | -34.7076 | -22.8332
14.0276 | -34.7081 | -22.8415
14.0337 | -34.7086 | -22.8499
14.0398 | -34.7091 | -22.8582
14.0459 | -34.7096 | -22.8665
14.052 | -34.71 | -22.875
14.0581 | -34.7105 | -22.8834
14.0642 | -34.7109 | -22.8917
14.0703 | -34.7114 | -22.9
14.0764 | -34.7119 | -22.9082
14.0825 | -34.7124 | -22.9165
14.0886 | -34.7129 | -22.9249
14.0947 | -34.7133 | -22.9332
14.1008 | -34.7138 | -22.9415
14.1069 | -34.7143 | -22.9497
14.1131 | -34.7147 | -22.9581
14.1192 | -34.7152 | -22.9663
14.1253 | -34.7157 | -22.9745
14.1314 | -34.7162 | -22.9828
14.1375 | -34.7167 | -22.9911
14.1436 | -34.7171 | -22.9994
14.1497 | -34.7176 | -23.0076
14.1558 | -34.7181 | -23.016
14.1619 | -34.7186 | -23.0242
14.168 | -34.719 | -23.0325
14.1741 | -34.7195 | -23.0407
14.1802 | -34.72 | -23.049
14.1863 | -34.7205 | -23.0573
14.1924 | -34.7209 | -23.0656
14.1985 | -34.7214 | -23.0738
14.2046 | -34.7219 | -23.0822
14.2107 | -34.7224 | -23.0902
14.2168 | -34.7229 | -23.0986
14.2229 | -34.7233 | -23.107
14.229 | -34.7238 | -23.1154
14.2351 | -34.7243 | -23.1237
14.2412 | -34.7247 | -23.132
14.2473 | -34.7252 | -23.1403
14.2534 | -34.7257 | -23.1485
14.2596 | -34.7262 | -23.1567
14.2657 | -34.7266 | -23.165
14.2718 | -34.7271 | -23.1733
14.2779 | -34.7276 | -23.1816
14.284 | -34.7281 | -23.19
14.2901 | -34.7285 | -23.1983
14.2962 | -34.729 | -23.2066
14.3023 | -34.7295 | -23.215
14.3084 | -34.73 | -23.2232
14.3145 | -34.7305 | -23.2315
14.3206 | -34.7309 | -23.24
14.3267 | -34.7314 | -23.2483
14.3328 | -34.7319 | -23.2566
14.3389 | -34.7324 | -23.2648
14.345 | -34.7328 | -23.2731
14.3511 | -34.7333 | -23.2813
14.3572 | -34.7338 | -23.2896
14.3633 | -34.7343 | -23.298
14.3694 | -34.7348 | -23.3062
14.3755 | -34.7352 | -23.3144
14.3816 | -34.7357 | -23.3228
14.3877 | -34.7362 | -23.331
14.3938 | -34.7367 | -23.3394
14.4 | -34.7372 | -23.3476
14.4061 | -34.7376 | -23.3559
14.4122 | -34.7381 | -23.3642
14.4183 | -34.7385 | -23.3724
14.4244 | -34.739 | -23.3806
14.4305 | -34.7395 | -23.3888
14.4366 | -34.74 | -23.3971
14.4427 | -34.7405 | -23.4055
14.4488 | -34.741 | -23.4139
14.4549 | -34.7414 | -23.422
14.461 | -34.7419 | -23.4301
14.4671 | -34.7424 | -23.4385
14.4732 | -34.7429 | -23.4468
14.4793 | -34.7433 | -23.4551
14.4854 | -34.7438 | -23.4633
14.4915 | -34.7443 | -23.4714
14.4976 | -34.7448 | -23.4796
14.5037 | -34.7453 | -23.4877
14.5098 | -34.7457 | -23.496
14.5159 | -34.7462 | -23.5042
14.522 | -34.7467 | -23.5124
14.5281 | -34.7471 | -23.5205
14.5342 | -34.7476 | -23.5288
14.5403 | -34.7481 | -23.537
14.5465 | -34.7486 | -23.5453
14.5526 | -34.7491 | -23.5537
14.5587 | -34.7495 | -23.562
14.5648 | -34.7501 | -23.5703
14.5709 | -34.7505 | -23.5785
14.577 | -34.751 | -23.5866
14.5831 | -34.7515 | -23.595
14.5892 | -34.7519 | -23.6031
14.5953 | -34.7524 | -23.6112
14.6014 | -34.7529 | -23.6195
14.6075 | -34.7534 | -23.6276
14.6136 | -34.7538 | -23.6358
14.6197 | -34.7543 | -23.644
14.6258 | -34.7547 | -23.6523
14.6319 | -34.7552 | -23.6604
14.638 | -34.7557 | -23.6687
14.6441 | -34.7562 | -23.677
14.6502 | -34.7566 | -23.6853
14.6563 | -34.7571 | -23.6936
14.6624 | -34.7576 | -23.7018
14.6685 | -34.7581 | -23.71
14.6746 | -34.7586 | -23.7182
14.6807 | -34.759 | -23.7263
14.6869 | -34.7595 | -23.7345
14.693 | -34.76 | -23.7426
14.6991 | -34.7605 | -23.7507
14.7052 | -34.761 | -23.7588
14.7113 | -34.7615 | -23.7672
14.7174 | -34.7619 | -23.7754
14.7235 | -34.7624 | -23.7835
14.7296 | -34.7629 | -23.7917
14.7357 | -34.7634 | -23.7999
14.7418 | -34.7639 | -23.8081
14.7479 | -34.7644 | -23.8162
14.754 | -34.7648 | -23.8244
14.7601 | -34.7653 | -23.8325
14.7662 | -34.7658 | -23.8405
14.7723 | -34.7663 | -23.8486
14.7784 | -34.7668 | -23.8569
14.7845 | -34.7672 | -23.8652
14.7906 | -34.7677 | -23.8735
14.7967 | -34.7682 | -23.8816
14.8028 | -34.7686 | -23.8897
14.8089 | -34.7691 | -23.8979
14.815 | -34.7696 | -23.9062
14.8211 | -34.7701 | -23.9144
14.8272 | -34.7706 | -23.9227
14.8334 | -34.7711 | -23.9309
14.8395 | -34.7715 | -23.9391
14.8456 | -34.772 | -23.9471
14.8517 | -34.7725 | -23.9553
14.8578 | -34.773 | -23.9635
14.8639 | -34.7734 | -23.9717
14.87 | -34.7739 | -23.98
14.8761 | -34.7744 | -23.988
14.8822 | -34.7749 | -23.9962
14.8883 | -34.7753 | -24.0044
14.8944 | -34.7758 | -24.0126
14.9005 | -34.7763 | -24.0208
14.9066 | -34.7768 | -24.0289
14.9127 | -34.7773 | -24.037
14.9188 | -34.7778 | -24.0453
14.9249 | -34.7782 | -24.0534
14.931 | -34.7787 | -24.0616
14.9371 | -34.7792 | -24.0696
14.9432 | -34.7797 | -24.0778
14.9493 | -34.7802 | -24.0858
14.9554 | -34.7806 | -24.0939
14.9615 | -34.7811 | -24.102
14.9676 | -34.7816 | -24.1102
14.9738 | -34.7821 | -24.1183
14.9799 | -34.7825 | -24.1267
14.986 | -34.783 | -24.1347
14.9921 | -34.7835 | -24.1428
14.9982 | -34.784 | -24.151
15.0043 | -34.7845 | -24.159
15.0104 | -34.785 | -24.1674
15.0165 | -34.7855 | -24.1756
15.0226 | -34.7859 | -24.1836
15.0287 | -34.7864 | -24.1918
15.0348 | -34.7869 | -24.2
15.0409 | -34.7873 | -24.2083
15.047 | -34.7878 | -24.2165
15.0531 | -34.7884 | -24.2245
15.0592 | -34.7888 | -24.2327
15.0653 | -34.7893 | -24.2409
15.0714 | -34.7898 | -24.249
15.0775 | -34.7903 | -24.2571
15.0836 | -34.7908 | -24.2652
15.0897 | -34.7913 | -24.2731
15.0958 | -34.7918 | -24.2812
15.1019 | -34.7923 | -24.2893
15.108 | -34.7928 | -24.2973
15.1141 | -34.7932 | -24.3054
15.1203 | -34.7937 | -24.3136
15.1264 | -34.7942 | -24.3216
15.1325 | -34.7947 | -24.3297
15.1386 | -34.7952 | -24.3379
15.1447 | -34.7957 | -24.3459
15.1508 | -34.7962 | -24.3541
15.1569 | -34.7967 | -24.3622
15.163 | -34.7972 | -24.37
15.1691 | -34.7976 | -24.3781
15.1752 | -34.7981 | -24.386
15.1813 | -34.7986 | -24.3942
15.1874 | -34.7991 | -24.4023
15.1935 | -34.7996 | -24.4103
15.1996 | -34.8 | -24.4183
15.2057 | -34.8005 | -24.4263
15.2118 | -34.801 | -24.4344
15.2179 | -34.8015 | -24.4426
15.224 | -34.802 | -24.4505
15.2301 | -34.8025 | -24.4584
15.2362 | -34.8029 | -24.4664
15.2423 | -34.8034 | -24.4746
15.2484 | -34.8039 | -24.4828
15.2545 | -34.8044 | -24.4909
15.2607 | -34.8049 | -24.499
15.2668 | -34.8054 | -24.5068
15.2729 | -34.8059 | -24.5148
15.279 | -34.8064 | -24.523
15.2851 | -34.8069 | -24.531
15.2912 | -34.8074 | -24.5391
15.2973 | -34.8079 | -24.5472
15.3034 | -34.8084 | -24.5552
15.3095 | -34.8089 | -24.5633
15.3156 | -34.8094 | -24.5712
15.3217 | -34.8099 | -24.5792
15.3278 | -34.8103 | -24.5874
15.3339 | -34.8109 | -24.5953
15.34 | -34.8113 | -24.6032
15.3461 | -34.8118 | -24.6111
15.3522 | -34.8123 | -24.6192
15.3583 | -34.8128 | -24.6274
15.3644 | -34.8133 | -24.6354
15.3705 | -34.8138 | -24.6433
15.3766 | -34.8143 | -24.6514
15.3827 | -34.8148 | -24.6594
15.3888 | -34.8153 | -24.6675
15.3949 | -34.8158 | -24.6756
15.401 | -34.8163 | -24.6837
15.4072 | -34.8168 | -24.6915
15.4133 | -34.8172 | -24.6996
15.4194 | -34.8177 | -24.7076
15.4255 | -34.8182 | -24.7158
15.4316 | -34.8187 | -24.7238
15.4377 | -34.8192 | -24.732
15.4438 | -34.8197 | -24.7401
15.4499 | -34.8202 | -24.7481
15.456 | -34.8207 | -24.756
15.4621 | -34.8212 | -24.7641
15.4682 | -34.8217 | -24.7723
15.4743 | -34.8222 | -24.7802
15.4804 | -34.8226 | -24.7881
15.4865 | -34.8231 | -24.796
15.4926 | -34.8236 | -24.8041
15.4987 | -34.8241 | -24.8121
15.5048 | -34.8246 | -24.8202
15.5109 | -34.8251 | -24.828
15.517 | -34.8256 | -24.836
15.5231 | -34.8261 | -24.844
15.5292 | -34.8266 | -24.8519
15.5353 | -34.8271 | -24.8601
15.5414 | -34.8276 | -24.8679
15.5476 | -34.8281 | -24.8757
15.5537 | -34.8285 | -24.8838
15.5598 | -34.8291 | -24.8918
15.5659 | -34.8295 | -24.8999
15.572 | -34.83 | -24.9078
15.5781 | -34.8305 | -24.9157
15.5842 | -34.831 | -24.9236
15.5903 | -34.8315 | -24.9315
15.5964 | -34.832 | -24.9394
15.6025 | -34.8325 | -24.9473
15.6086 | -34.833 | -24.9553
15.6147 | -34.8335 | -24.9634
15.6208 | -34.834 | -24.9712
15.6269 | -34.8345 | -24.9791
15.633 | -34.835 | -24.9871
15.6391 | -34.8355 | -24.9951
15.6452 | -34.836 | -25.0031
15.6513 | -34.8365 | -25.0109
15.6574 | -34.837 | -25.0189
15.6635 | -34.8375 | -25.0269
15.6696 | -34.838 | -25.0349
15.6757 | -34.8385 | -25.043
15.6818 | -34.839 | -25.0508
15.688 | -34.8395 | -25.0585
15.6941 | -34.84 | -25.0663
15.7002 | -34.8404 | -25.0742
15.7063 | -34.8409 | -25.082
15.7124 | -34.8414 | -25.0899
15.7185 | -34.8419 | -25.0979
15.7246 | -34.8424 | -25.1058
15.7307 | -34.843 | -25.1138
15.7368 | -34.8435 | -25.1217
15.7429 | -34.844 | -25.1296
15.749 | -34.8445 | -25.1377
15.7551 | -34.8449 | -25.1457
15.7612 | -34.8455 | -25.1534
15.7673 | -34.846 | -25.1614
15.7734 | -34.8465 | -25.1695
15.7795 | -34.847 | -25.1773
15.7856 | -34.8475 | -25.1852
15.7917 | -34.848 | -25.193
15.7978 | -34.8485 | -25.2011
15.8039 | -34.849 | -25.209
15.81 | -34.8495 | -25.2169
15.8161 | -34.8499 | -25.2248
15.8222 | -34.8504 | -25.2326
15.8283 | -34.8509 | -25.2404
15.8345 | -34.8514 | -25.2482
15.8406 | -34.8519 | -25.256
15.8467 | -34.8524 | -25.2639
15.8528 | -34.8529 | -25.2719
15.8589 | -34.8534 | -25.2798
15.865 | -34.854 | -25.2875
15.8711 | -34.8545 | -25.2956
15.8772 | -34.855 | -25.3035
15.8833 | -34.8555 | -25.3115
15.8894 | -34.856 | -25.3193
15.8955 | -34.8565 | -25.3272
15.9016 | -34.857 | -25.3348
15.9077 | -34.8575 | -25.3428
15.9138 | -34.8579 | -25.3507
15.9199 | -34.8584 | -25.3585
15.926 | -34.8589 | -25.3662
15.9321 | -34.8594 | -25.374
15.9382 | -34.8599 | -25.3817
15.9443 | -34.8604 | -25.3897
15.9504 | -34.861 | -25.3974
15.9565 | -34.8614 | -25.4053
15.9626 | -34.8619 | -25.4133
15.9687 | -34.8624 | -25.4211
15.9749 | -34.8629 | -25.4288
15.981 | -34.8634 | -25.4366
15.9871 | -34.8639 | -25.4445
15.9932 | -34.8644 | -25.4523
15.9993 | -34.8649 | -25.4601
16.0054 | -34.8654 | -25.4676
16.0115 | -34.8659 | -25.4755
16.0176 | -34.8664 | -25.4833
16.0237 | -34.8669 | -25.4913
16.0298 | -34.8673 | -25.4991
16.0359 | -34.8678 | -25.5069
16.042 | -34.8684 | -25.5147
16.0481 | -34.8689 | -25.5225
16.0542 | -34.8694 | -25.5304
16.0603 | -34.8698 | -25.5382
16.0664 | -34.8704 | -25.546
16.0725 | -34.8709 | -25.5536
16.0786 | -34.8713 | -25.5614
16.0847 | -34.8718 | -25.5693
16.0908 | -34.8723 | -25.5773
16.0969 | -34.8728 | -25.5851
16.103 | -34.8733 | -25.5929
16.1091 | -34.8738 | -25.6007
16.1152 | -34.8743 | -25.6084
16.1214 | -34.8748 | -25.6162
16.1275 | -34.8753 | -25.6241
16.1336 | -34.8758 | -25.6318
16.1397 | -34.8763 | -25.6397
16.1458 | -34.8768 | -25.6474
16.1519 | -34.8773 | -25.6553
16.158 | -34.8777 | -25.6634
16.1641 | -34.8782 | -25.6714
16.1702 | -34.8787 | -25.6792
16.1763 | -34.8792 | -25.687
16.1824 | -34.8797 | -25.6946
16.1885 | -34.8802 | -25.7023
16.1946 | -34.8807 | -25.7101
16.2007 | -34.8812 | -25.718
16.2068 | -34.8817 | -25.7258
16.2129 | -34.8822 | -25.7336
16.219 | -34.8827 | -25.741
16.2251 | -34.8832 | -25.7486
16.2312 | -34.8837 | -25.7564
16.2373 | -34.8842 | -25.764
16.2434 | -34.8847 | -25.7718
16.2495 | -34.8852 | -25.7793
16.2556 | -34.8857 | -25.7873
16.2618 | -34.8862 | -25.7949
16.2679 | -34.8867 | -25.8028
16.274 | -34.8873 | -25.8106
16.2801 | -34.8877 | -25.8184
16.2862 | -34.8882 | -25.8261
16.2923 | -34.8888 | -25.8338
16.2984 | -34.8893 | -25.8415
16.3045 | -34.8897 | -25.8492
16.3106 | -34.8902 | -25.857
16.3167 | -34.8907 | -25.8649
16.3228 | -34.8912 | -25.8725
16.3289 | -34.8917 | -25.8803
16.335 | -34.8922 | -25.888
16.3411 | -34.8927 | -25.8958
16.3472 | -34.8932 | -25.9035
16.3533 | -34.8937 | -25.9113
16.3594 | -34.8942 | -25.9189
16.3655 | -34.8947 | -25.9267
16.3716 | -34.8952 | -25.9344
16.3777 | -34.8957 | -25.942
16.3838 | -34.8962 | -25.9498
16.3899 | -34.8967 | -25.9575
16.396 | -34.8972 | -25.9653
16.4021 | -34.8977 | -25.973
16.4083 | -34.8982 | -25.9807
16.4144 | -34.8987 | -25.9885
16.4205 | -34.8992 | -25.9961
16.4266 | -34.8997 | -26.0037
16.4327 | -34.9002 | -26.0114
16.4388 | -34.9007 | -26.0191
16.4449 | -34.9012 | -26.0268
16.451 | -34.9017 | -26.0345
16.4571 | -34.9022 | -26.0425
16.4632 | -34.9027 | -26.0501
16.4693 | -34.9031 | -26.0578
16.4754 | -34.9037 | -26.0656
16.4815 | -34.9042 | -26.0735
16.4876 | -34.9047 | -26.0812
16.4937 | -34.9051 | -26.0888
16.4998 | -34.9057 | -26.0963
16.5059 | -34.9062 | -26.1038
16.512 | -34.9067 | -26.1117
16.5181 | -34.9072 | -26.1194
16.5242 | -34.9077 | -26.1274
16.5303 | -34.9082 | -26.1351
16.5364 | -34.9087 | -26.1427
16.5425 | -34.9092 | -26.1506
16.5487 | -34.9096 | -26.1583
16.5548 | -34.9101 | -26.166
16.5609 | -34.9106 | -26.1736
16.567 | -34.9111 | -26.1811
16.5731 | -34.9117 | -26.1889
16.5792 | -34.9122 | -26.1967
16.5853 | -34.9127 | -26.2045
16.5914 | -34.9132 | -26.2123
16.5975 | -34.9137 | -26.2201
16.6036 | -34.9142 | -26.2277
16.6097 | -34.9147 | -26.2353
16.6158 | -34.9152 | -26.2429
16.6219 | -34.9157 | -26.2507
16.628 | -34.9162 | -26.2584
16.6341 | -34.9167 | -26.266
16.6402 | -34.9172 | -26.2735
16.6463 | -34.9177 | -26.2812
16.6524 | -34.9181 | -26.2888
16.6585 | -34.9186 | -26.2965
16.6646 | -34.9191 | -26.3041
16.6707 | -34.9196 | -26.3116
16.6768 | -34.9201 | -26.3193
16.6829 | -34.9206 | -26.327
16.689 | -34.9211 | -26.3348
16.6952 | -34.9216 | -26.3425
16.7013 | -34.9221 | -26.3501
16.7074 | -34.9226 | -26.3575
16.7135 | -34.9231 | -26.3653
16.7196 | -34.9236 | -26.3729
16.7257 | -34.9242 | -26.3806
16.7318 | -34.9246 | -26.3882
16.7379 | -34.9251 | -26.3958
16.744 | -34.9256 | -26.4033
16.7501 | -34.9261 | -26.4112
16.7562 | -34.9266 | -26.4188
16.7623 | -34.9271 | -26.4263
16.7684 | -34.9276 | -26.4341
16.7745 | -34.928 | -26.442
16.7806 | -34.9285 | -26.4496
16.7867 | -34.9291 | -26.4572
16.7928 | -34.9295 | -26.4648
16.7989 | -34.93 | -26.4724
16.805 | -34.9305 | -26.4802
16.8111 | -34.931 | -26.4879
16.8172 | -34.9315 | -26.4956
16.8233 | -34.932 | -26.5034
16.8294 | -34.9325 | -26.511
16.8356 | -34.933 | -26.5187
16.8417 | -34.9335 | -26.5264
16.8478 | -34.9341 | -26.5339
16.8539 | -34.9346 | -26.5414
16.86 | -34.9351 | -26.5492
16.8661 | -34.9356 | -26.5567
16.8722 | -34.9361 | -26.5644
16.8783 | -34.9366 | -26.5722
16.8844 | -34.9371 | -26.5797
16.8905 | -34.9376 | -26.5873
16.8966 | -34.9381 | -26.5948
16.9027 | -34.9386 | -26.6024
16.9088 | -34.9391 | -26.6102
16.9149 | -34.9396 | -26.618
16.921 | -34.9401 | -26.6251
16.9271 | -34.9406 | -26.6327
16.9332 | -34.9411 | -26.6402
16.9393 | -34.9416 | -26.6478
16.9454 | -34.9421 | -26.6553
16.9515 | -34.9426 | -26.6631
16.9576 | -34.9431 | -26.6704
16.9637 | -34.9437 | -26.6779
16.9698 | -34.9442 | -26.6855
16.9759 | -34.9447 | -26.693
16.9821 | -34.9452 | -26.7006
16.9882 | -34.9457 | -26.7078
16.9943 | -34.9461 | -26.7152
17.0004 | -34.9466 | -26.7226
17.0065 | -34.9472 | -26.7303
17.0126 | -34.9477 | -26.7378
17.0187 | -34.9482 | -26.7453
17.0248 | -34.9487 | -26.7528
17.0309 | -34.9492 | -26.7602
17.037 | -34.9497 | -26.7678
17.0431 | -34.9502 | -26.7753
17.0492 | -34.9507 | -26.7831
17.0553 | -34.9512 | -26.7906
17.0614 | -34.9516 | -26.7981
17.0675 | -34.9521 | -26.8055
17.0736 | -34.9526 | -26.8133
17.0797 | -34.9531 | -26.8205
17.0858 | -34.9536 | -26.8282
17.0919 | -34.9541 | -26.8357
17.098 | -34.9546 | -26.8431
17.1041 | -34.9551 | -26.8506
17.1102 | -34.9556 | -26.8581
17.1163 | -34.9561 | -26.8655
17.1225 | -34.9566 | -26.8731
17.1286 | -34.9572 | -26.8803
17.1347 | -34.9577 | -26.8879
17.1408 | -34.9581 | -26.8956
17.1469 | -34.9586 | -26.9031
17.153 | -34.9591 | -26.9105
17.1591 | -34.9596 | -26.918
17.1652 | -34.9601 | -26.9251
17.1713 | -34.9606 | -26.9326
17.1774 | -34.9611 | -26.9401
17.1835 | -34.9616 | -26.9476
17.1896 | -34.9621 | -26.9551
17.1957 | -34.9626 | -26.9627
17.2018 | -34.9631 | -26.9698
17.2079 | -34.9636 | -26.9775
17.214 | -34.9641 | -26.9851
17.2201 | -34.9646 | -26.9926
17.2262 | -34.9651 | -27.0002
17.2323 | -34.9656 | -27.0078
17.2384 | -34.9661 | -27.0154
17.2445 | -34.9665 | -27.0228
17.2506 | -34.967 | -27.0303
17.2567 | -34.9675 | -27.0378
17.2628 | -34.9681 | -27.0454
17.269 | -34.9685 | -27.0528
17.2751 | -34.969 | -27.0604
17.2812 | -34.9695 | -27.0682
17.2873 | -34.97 | -27.0757
17.2934 | -34.9705 | -27.0833
17.2995 | -34.971 | -27.091
17.3056 | -34.9715 | -27.0983
17.3117 | -34.972 | -27.1058
17.3178 | -34.9725 | -27.1133
17.3239 | -34.973 | -27.1209
17.33 | -34.9735 | -27.1283
17.3361 | -34.974 | -27.1358
17.3422 | -34.9745 | -27.1429
17.3483 | -34.975 | -27.1503
17.3544 | -34.9755 | -27.1578
17.3605 | -34.976 | -27.1653
17.3666 | -34.9764 | -27.1728
17.3727 | -34.9769 | -27.1803
17.3788 | -34.9774 | -27.1874
17.3849 | -34.978 | -27.1949
17.391 | -34.9784 | -27.2023
17.3971 | -34.9789 | -27.2098
17.4032 | -34.9794 | -27.2173
17.4094 | -34.9799 | -27.2245
17.4155 | -34.9804 | -27.2318
17.4216 | -34.9809 | -27.2392
17.4277 | -34.9814 | -27.2468
17.4338 | -34.9819 | -27.2542
17.4399 | -34.9824 | -27.2617
17.446 | -34.9829 | -27.269
17.4521 | -34.9834 | -27.2764
17.4582 | -34.9839 | -27.2838
17.4643 | -34.9844 | -27.2912
17.4704 | -34.9849 | -27.2986
17.4765 | -34.9854 | -27.3061
17.4826 | -34.9859 | -27.3134
17.4887 | -34.9864 | -27.3209
17.4948 | -34.9869 | -27.3284
17.5009 | -34.9874 | -27.3357
17.507 | -34.9879 | -27.3432
17.5131 | -34.9884 | -27.3506
17.5192 | -34.9889 | -27.3577
17.5253 | -34.9893 | -27.3651
17.5314 | -34.9898 | -27.3725
17.5375 | -34.9903 | -27.38
17.5436 | -34.9908 | -27.3875
17.5497 | -34.9913 | -27.3949
17.5559 | -34.9918 | -27.4023
17.562 | -34.9923 | -27.4097
17.5681 | -34.9928 | -27.417
17.5742 | -34.9933 | -27.4246
17.5803 | -34.9938 | -27.4319
17.5864 | -34.9943 | -27.4392
17.5925 | -34.9948 | -27.4467
17.5986 | -34.9953 | -27.454
17.6047 | -34.9958 | -27.4615
17.6108 | -34.9963 | -27.4687
17.6169 | -34.9968 | -27.4761
17.623 | -34.9973 | -27.4833
17.6291 | -34.9978 | -27.4906
17.6352 | -34.9983 | -27.4979
17.6413 | -34.9988 | -27.5053
17.6474 | -34.9993 | -27.5128
17.6535 | -34.9998 | -27.5201
17.6596 | -35.0003 | -27.5272
17.6657 | -35.0008 | -27.5347
17.6718 | -35.0013 | -27.542
17.6779 | -35.0017 | -27.5492
17.684 | -35.0023 | -27.5566
17.6901 | -35.0027 | -27.5639
17.6963 | -35.0032 | -27.571
17.7024 | -35.0037 | -27.5782
17.7085 | -35.0042 | -27.5855
17.7146 | -35.0047 | -27.5928
17.7207 | -35.0052 | -27.6
17.7268 | -35.0057 | -27.6072
17.7329 | -35.0062 | -27.6146
17.739 | -35.0067 | -27.622
17.7451 | -35.0072 | -27.6293
17.7512 | -35.0077 | -27.6368
17.7573 | -35.0082 | -27.6442
17.7634 | -35.0087 | -27.6513
17.7695 | -35.0092 | -27.6586
17.7756 | -35.0097 | -27.666
17.7817 | -35.0102 | -27.6733
17.7878 | -35.0107 | -27.6806
17.7939 | -35.0111 | -27.688
17.8 | -35.0116 | -27.6952
17.8061 | -35.0122 | -27.7026
17.8122 | -35.0127 | -27.7099
17.8183 | -35.0132 | -27.7172
17.8244 | -35.0137 | -27.7247
17.8305 | -35.0142 | -27.732
17.8366 | -35.0147 | -27.7391
17.8428 | -35.0152 | -27.7464
17.8489 | -35.0157 | -27.7537
17.855 | -35.0162 | -27.7609
17.8611 | -35.0166 | -27.7682
17.8672 | -35.0171 | -27.7752
17.8733 | -35.0176 | -27.7826
17.8794 | -35.0181 | -27.7899
17.8855 | -35.0186 | -27.7972
17.8916 | -35.0191 | -27.8047
17.8977 | -35.0196 | -27.8121
17.9038 | -35.0201 | -27.8191
17.9099 | -35.0205 | -27.8264
17.916 | -35.021 | -27.8338
17.9221 | -35.0215 | -27.841
17.9282 | -35.022 | -27.8483
17.9343 | -35.0225 | -27.8555
17.9404 | -35.023 | -27.8626
17.9465 | -35.0235 | -27.8699
17.9526 | -35.024 | -27.8773
17.9587 | -35.0245 | -27.8845
17.9648 | -35.025 | -27.8917
17.9709 | -35.0255 | -27.8989
17.977 | -35.026 | -27.9057
17.9832 | -35.0265 | -27.9129
17.9893 | -35.027 | -27.9201
17.9954 | -35.0274 | -27.9273
18.0015 | -35.0279 | -27.9346
18.0076 | -35.0284 | -27.9417
18.0137 | -35.0289 | -27.949
18.0198 | -35.0294 | -27.9563
18.0259 | -35.0299 | -27.9637
18.032 | -35.0304 | -27.9709
18.0381 | -35.0309 | -27.9781
18.0442 | -35.0314 | -27.9853
18.0503 | -35.0319 | -27.9926
18.0564 | -35.0324 | -27.9998
18.0625 | -35.0329 | -28.007
18.0686 | -35.0333 | -28.0143
18.0747 | -35.0338 | -28.0217
18.0808 | -35.0343 | -28.0287
18.0869 | -35.0348 | -28.0358
18.093 | -35.0353 | -28.0431
18.0991 | -35.0358 | -28.0503
18.1052 | -35.0363 | -28.0576
18.1113 | -35.0368 | -28.0647
18.1174 | -35.0372 | -28.0717
18.1236 | -35.0377 | -28.079
18.1297 | -35.0382 | -28.0862
18.1358 | -35.0387 | -28.0935
18.1419 | -35.0392 | -28.1007
18.148 | -35.0397 | -28.1079
18.1541 | -35.0402 | -28.1147
18.1602 | -35.0407 | -28.1219
18.1663 | -35.0412 | -28.1291
18.1724 | -35.0416 | -28.1363
18.1785 | -35.0421 | -28.1435
18.1846 | -35.0426 | -28.1506
18.1907 | -35.0431 | -28.1577
18.1968 | -35.0436 | -28.165
18.2029 | -35.0441 | -28.1722
18.209 | -35.0446 | -28.1795
18.2151 | -35.0451 | -28.1869
18.2212 | -35.0456 | -28.1939
18.2273 | -35.0461 | -28.2013
18.2334 | -35.0466 | -28.2085
18.2395 | -35.047 | -28.2159
18.2456 | -35.0475 | -28.2231
18.2517 | -35.048 | -28.2304
18.2578 | -35.0485 | -28.2376
18.2639 | -35.049 | -28.2448
18.2701 | -35.0495 | -28.252
18.2762 | -35.05 | -28.2591
18.2823 | -35.0505 | -28.2663
18.2884 | -35.051 | -28.2734
18.2945 | -35.0515 | -28.2802
18.3006 | -35.052 | -28.2874
18.3067 | -35.0524 | -28.2946
18.3128 | -35.0529 | -28.3019
18.3189 | -35.0534 | -28.309
18.325 | -35.0539 | -28.316
18.3311 | -35.0544 | -28.3232
18.3372 | -35.0549 | -28.3304
18.3433 | -35.0554 | -28.3377
18.3494 | -35.0558 | -28.3449
18.3555 | -35.0563 | -28.352
18.3616 | -35.0568 | -28.3588
18.3677 | -35.0573 | -28.3661
18.3738 | -35.0578 | -28.3733
18.3799 | -35.0583 | -28.3805
18.386 | -35.0587 | -28.3876
18.3921 | -35.0592 | -28.3947
18.3982 | -35.0597 | -28.4014
18.4043 | -35.0602 | -28.4086
18.4105 | -35.0607 | -28.4158
18.4166 | -35.0611 | -28.423
18.4227 | -35.0616 | -28.4302
18.4288 | -35.0621 | -28.4374
18.4349 | -35.0626 | -28.4443
18.441 | -35.0631 | -28.4514
18.4471 | -35.0636 | -28.4584
18.4532 | -35.064 | -28.4656
18.4593 | -35.0645 | -28.4728
18.4654 | -35.065 | -28.4796
18.4715 | -35.0655 | -28.4868
18.4776 | -35.066 | -28.4941
18.4837 | -35.0665 | -28.5012
18.4898 | -35.067 | -28.5084
18.4959 | -35.0675 | -28.5157
18.502 | -35.0679 | -28.5226
18.5081 | -35.0684 | -28.5297
18.5142 | -35.0689 | -28.5367
18.5203 | -35.0694 | -28.5438
18.5264 | -35.0699 | -28.5509
18.5325 | -35.0704 | -28.558
18.5386 | -35.0709 | -28.5649
18.5447 | -35.0714 | -28.5719
18.5508 | -35.0718 | -28.5791
18.557 | -35.0723 | -28.5861
18.5631 | -35.0728 | -28.5932
18.5692 | -35.0733 | -28.6004
18.5753 | -35.0737 | -28.6073
18.5814 | -35.0742 | -28.6145
18.5875 | -35.0747 | -28.6217
18.5936 | -35.0752 | -28.6287
18.5997 | -35.0756 | -28.6359
18.6058 | -35.0761 | -28.6429
18.6119 | -35.0766 | -28.6499
18.618 | -35.0771 | -28.657
18.6241 | -35.0776 | -28.6641
18.6302 | -35.078 | -28.6713
18.6363 | -35.0785 | -28.6784
18.6424 | -35.079 | -28.6853
18.6485 | -35.0795 | -28.6923
18.6546 | -35.0799 | -28.6994
18.6607 | -35.0804 | -28.7062
18.6668 | -35.0809 | -28.7134
18.6729 | -35.0814 | -28.7205
18.679 | -35.0819 | -28.7272
18.6851 | -35.0824 | -28.7342
18.6912 | -35.0828 | -28.7413
18.6974 | -35.0833 | -28.7483
18.7035 | -35.0838 | -28.7554
18.7096 | -35.0843 | -28.7624
18.7157 | -35.0848 | -28.7691
18.7218 | -35.0852 | -28.7761
18.7279 | -35.0858 | -28.783
18.734 | -35.0862 | -28.7902
18.7401 | -35.0867 | -28.7972
18.7462 | -35.0872 | -28.8038
18.7523 | -35.0877 | -28.811
18.7584 | -35.0882 | -28.8181
18.7645 | -35.0887 | -28.825
18.7706 | -35.0892 | -28.8322
18.7767 | -35.0896 | -28.8391
18.7828 | -35.0901 | -28.8459
18.7889 | -35.0906 | -28.8528
18.795 | -35.0911 | -28.8597
18.8011 | -35.0916 | -28.8667
18.8072 | -35.0921 | -28.8738
18.8133 | -35.0926 | -28.8809
18.8194 | -35.0931 | -28.8878
18.8255 | -35.0936 | -28.8947
18.8316 | -35.094 | -28.9017
18.8377 | -35.0945 | -28.9085
18.8439 | -35.095 | -28.9156
18.85 | -35.0955 | -28.9227
18.8561 | -35.0959 | -28.9298
18.8622 | -35.0964 | -28.9367
18.8683 | -35.0969 | -28.9439
18.8744 | -35.0974 | -28.9509
18.8805 | -35.0979 | -28.9579
18.8866 | -35.0983 | -28.9649
18.8927 | -35.0989 | -28.9716
18.8988 | -35.0993 | -28.9787
18.9049 | -35.0998 | -28.9856
18.911 | -35.1003 | -28.9926
18.9171 | -35.1008 | -28.9996
18.9232 | -35.1013 | -29.0063
18.9293 | -35.1018 | -29.0133
18.9354 | -35.1023 | -29.0203
18.9415 | -35.1028 | -29.0273
18.9476 | -35.1032 | -29.0343
18.9537 | -35.1037 | -29.0414
18.9598 | -35.1042 | -29.0481
18.9659 | -35.1047 | -29.0551
18.972 | -35.1051 | -29.0621
18.9781 | -35.1056 | -29.0691
18.9843 | -35.106 | -29.0762
18.9904 | -35.1065 | -29.0832
18.9965 | -35.1069 | -29.0898
19.0026 | -35.1074 | -29.0966
19.0087 | -35.1079 | -29.1034
19.0148 | -35.1084 | -29.1104
19.0209 | -35.1089 | -29.1174
19.027 | -35.1094 | -29.1242
19.0331 | -35.1098 | -29.1309
19.0392 | -35.1103 | -29.1381
19.0453 | -35.1108 | -29.145
19.0514 | -35.1113 | -29.1518
19.0575 | -35.1118 | -29.1588
19.0636 | -35.1122 | -29.1651
19.0697 | -35.1127 | -29.1722
19.0758 | -35.1132 | -29.1792
19.0819 | -35.1137 | -29.186
19.088 | -35.1142 | -29.1931
19.0941 | -35.1146 | -29.2001
19.1002 | -35.1151 | -29.2068
19.1063 | -35.1155 | -29.2138
19.1124 | -35.116 | -29.2208
19.1185 | -35.1165 | -29.2277
19.1246 | -35.117 | -29.2344
19.1308 | -35.1175 | -29.2414
19.1369 | -35.118 | -29.2481
19.143 | -35.1185 | -29.255
19.1491 | -35.119 | -29.2621
19.1552 | -35.1194 | -29.269
19.1613 | -35.1199 | -29.276
19.1674 | -35.1204 | -29.2829
19.1735 | -35.1208 | -29.2897
19.1796 | -35.1213 | -29.2966
19.1857 | -35.1218 | -29.3036
19.1918 | -35.1223 | -29.3106
19.1979 | -35.1228 | -29.3174
19.204 | -35.1232 | -29.3241
19.2101 | -35.1237 | -29.3311
19.2162 | -35.1241 | -29.3379
19.2223 | -35.1246 | -29.3447
19.2284 | -35.1251 | -29.3516
19.2345 | -35.1256 | -29.3584
19.2406 | -35.1261 | -29.3652
19.2467 | -35.1265 | -29.3721
19.2528 | -35.127 | -29.3792
19.2589 | -35.1275 | -29.386
19.265 | -35.128 | -29.3929
19.2712 | -35.1284 | -29.3998
19.2773 | -35.1289 | -29.4066
19.2834 | -35.1294 | -29.4134
19.2895 | -35.1299 | -29.4203
19.2956 | -35.1303 | -29.4271
19.3017 | -35.1308 | -29.4339
19.3078 | -35.1313 | -29.4407
19.3139 | -35.1318 | -29.4474
19.32 | -35.1323 | -29.4542
19.3261 | -35.1327 | -29.4609
19.3322 | -35.1332 | -29.468
19.3383 | -35.1337 | -29.4747
19.3444 | -35.1342 | -29.4813
19.3505 | -35.1347 | -29.4882
19.3566 | -35.1351 | -29.4952
19.3627 | -35.1356 | -29.5022
19.3688 | -35.1361 | -29.509
19.3749 | -35.1366 | -29.5158
19.381 | -35.1371 | -29.5226
19.3871 | -35.1375 | -29.5295
19.3932 | -35.138 | -29.5366
19.3993 | -35.1385 | -29.5435
19.4054 | -35.139 | -29.5506
19.4115 | -35.1395 | -29.5575
19.4177 | -35.1399 | -29.5643
19.4238 | -35.1404 | -29.5711
19.4299 | -35.1409 | -29.5779
19.436 | -35.1414 | -29.5849
19.4421 | -35.1419 | -29.5919
19.4482 | -35.1423 | -29.5989
19.4543 | -35.1428 | -29.6054
19.4604 | -35.1432 | -29.6122
19.4665 | -35.1437 | -29.6191
19.4726 | -35.1441 | -29.6257
19.4787 | -35.1446 | -29.6326
19.4848 | -35.145 | -29.6395
19.4909 | -35.1455 | -29.6462
19.497 | -35.146 | -29.653
19.5031 | -35.1464 | -29.6597
19.5092 | -35.1469 | -29.6665
19.5153 | -35.1474 | -29.6732
19.5214 | -35.1478 | -29.6797
19.5275 | -35.1483 | -29.6865
19.5336 | -35.1487 | -29.6932
19.5397 | -35.1492 | -29.7001
19.5458 | -35.1497 | -29.7069
19.5519 | -35.1502 | -29.7137
19.5581 | -35.1507 | -29.7204
19.5642 | -35.1511 | -29.7273
19.5703 | -35.1516 | -29.7341
19.5764 | -35.152 | -29.7408
19.5825 | -35.1525 | -29.7473
19.5886 | -35.153 | -29.7541
19.5947 | -35.1535 | -29.761
19.6008 | -35.1539 | -29.7679
19.6069 | -35.1543 | -29.7748
19.613 | -35.1548 | -29.7818
19.6191 | -35.1553 | -29.7886
19.6252 | -35.1558 | -29.7954
19.6313 | -35.1562 | -29.8019
19.6374 | -35.1566 | -29.8085
19.6435 | -35.1571 | -29.8152
19.6496 | -35.1576 | -29.8221
19.6557 | -35.1581 | -29.829
19.6618 | -35.1585 | -29.8359
19.6679 | -35.159 | -29.8427
19.674 | -35.1594 | -29.8496
19.6801 | -35.1599 | -29.8565
19.6862 | -35.1604 | -29.8632
19.6923 | -35.1608 | -29.8697
19.6984 | -35.1613 | -29.8764
19.7046 | -35.1618 | -29.8832
19.7107 | -35.1623 | -29.8899
19.7168 | -35.1627 | -29.8967
19.7229 | -35.1632 | -29.9037
19.729 | -35.1637 | -29.9104
19.7351 | -35.1641 | -29.9166
19.7412 | -35.1646 | -29.9234
19.7473 | -35.1651 | -29.9302
19.7534 | -35.1655 | -29.937
19.7595 | -35.1659 | -29.9437
19.7656 | -35.1664 | -29.9508
19.7717 | -35.1668 | -29.9573
19.7778 | -35.1673 | -29.9641
19.7839 | -35.1677 | -29.9707
19.79 | -35.1682 | -29.9774
19.7961 | -35.1686 | -29.9844
19.8022 | -35.1691 | -29.9912
19.8083 | -35.1696 | -29.9978
19.8144 | -35.17 | -30.0047
19.8205 | -35.1705 | -30.0114
19.8266 | -35.171 | -30.0181
19.8327 | -35.1714 | -30.0251
19.8388 | -35.1719 | -30.0319
19.845 | -35.1723 | -30.0386
19.8511 | -35.1727 | -30.0456
19.8572 | -35.1732 | -30.0524
19.8633 | -35.1737 | -30.0592
19.8694 | -35.1741 | -30.0662
19.8755 | -35.1746 | -30.073
19.8816 | -35.175 | -30.0799
19.8877 | -35.1755 | -30.0868
19.8938 | -35.1759 | -30.0935
19.8999 | -35.1764 | -30.1004
19.906 | -35.1769 | -30.1073
19.9121 | -35.1774 | -30.1138
19.9182 | -35.1778 | -30.1205
19.9243 | -35.1783 | -30.1272
19.9304 | -35.1788 | -30.1341
19.9365 | -35.1793 | -30.1408
19.9426 | -35.1797 | -30.1475
19.9487 | -35.1802 | -30.1543
19.9548 | -35.1806 | -30.1611
19.9609 | -35.1811 | -30.1679
19.967 | -35.1815 | -30.1747
19.9731 | -35.182 | -30.1814
19.9792 | -35.1825 | -30.188
19.9853 | -35.1829 | -30.1944
19.9915 | -35.1834 | -30.2011
19.9976 | -35.1838 | -30.2078
20.0037 | -35.1843 | -30.2145
20.0098 | -35.1847 | -30.2213
20.0159 | -35.1852 | -30.2281
20.022 | -35.1856 | -30.2348
20.0281 | -35.1861 | -30.2418
20.0342 | -35.1865 | -30.2484
20.0403 | -35.187 | -30.2551
20.0464 | -35.1875 | -30.2618
20.0525 | -35.188 | -30.2684
20.0586 | -35.1885 | -30.2751
20.0647 | -35.1889 | -30.282
20.0708 | -35.1894 | -30.2889
20.0769 | -35.1899 | -30.2956
20.083 | -35.1903 | -30.3024
20.0891 | -35.1908 | -30.3094
20.0952 | -35.1912 | -30.3163
20.1013 | -35.1917 | -30.3231
20.1074 | -35.1921 | -30.3297
20.1135 | -35.1926 | -30.3363
20.1196 | -35.193 | -30.3429
20.1257 | -35.1935 | -30.3498
20.1319 | -35.1939 | -30.3568
20.138 | -35.1944 | -30.3635
20.1441 | -35.1949 | -30.3704
20.1502 | -35.1953 | -30.3773
20.1563 | -35.1958 | -30.3839
20.1624 | -35.1962 | -30.3905
20.1685 | -35.1966 | -30.3973
20.1746 | -35.1971 | -30.4039
20.1807 | -35.1976 | -30.4106
20.1868 | -35.1981 | -30.4173
20.1929 | -35.1985 | -30.424
20.199 | -35.199 | -30.4309
20.2051 | -35.1994 | -30.4376
20.2112 | -35.1999 | -30.4443
20.2173 | -35.2003 | -30.4509
20.2234 | -35.2007 | -30.4576
20.2295 | -35.2012 | -30.4645
20.2356 | -35.2017 | -30.4714
20.2417 | -35.2022 | -30.4782
20.2478 | -35.2026 | -30.485
20.2539 | -35.2031 | -30.492
20.26 | -35.2035 | -30.4989
20.2661 | -35.204 | -30.5056
20.2723 | -35.2045 | -30.5124
20.2784 | -35.205 | -30.5189
20.2845 | -35.2054 | -30.5256
20.2906 | -35.2059 | -30.5321
20.2967 | -35.2063 | -30.5388
20.3028 | -35.2068 | -30.5455
20.3089 | -35.2072 | -30.5522
20.315 | -35.2077 | -30.5588
20.3211 | -35.2081 | -30.5656
20.3272 | -35.2086 | -30.5725
20.3333 | -35.209 | -30.5792
20.3394 | -35.2094 | -30.586
20.3455 | -35.2099 | -30.5927
20.3516 | -35.2103 | -30.5994
20.3577 | -35.2107 | -30.6062
20.3638 | -35.2112 | -30.6129
20.3699 | -35.2117 | -30.6193
20.376 | -35.2121 | -30.6259
20.3821 | -35.2126 | -30.6327
20.3882 | -35.213 | -30.6395
20.3943 | -35.2135 | -30.6463
20.4004 | -35.2139 | -30.6528
20.4065 | -35.2144 | -30.6596
20.4126 | -35.2148 | -30.6663
20.4188 | -35.2153 | -30.6732
20.4249 | -35.2158 | -30.6797
20.431 | -35.2162 | -30.6863
20.4371 | -35.2166 | -30.693
20.4432 | -35.2171 | -30.6997
20.4493 | -35.2176 | -30.7065
20.4554 | -35.218 | -30.7132
20.4615 | -35.2185 | -30.72
20.4676 | -35.219 | -30.727
20.4737 | -35.2194 | -30.7336
20.4798 | -35.2199 | -30.7403
20.4859 | -35.2203 | -30.7471
20.492 | -35.2208 | -30.7536
20.4981 | -35.2212 | -30.7601
20.5042 | -35.2216 | -30.7667
20.5103 | -35.2221 | -30.7732
20.5164 | -35.2225 | -30.7799
20.5225 | -35.223 | -30.7867
20.5286 | -35.2234 | -30.7933
20.5347 | -35.2238 | -30.7999
20.5408 | -35.2243 | -30.8067
20.5469 | -35.2247 | -30.8137
20.553 | -35.2252 | -30.8205
20.5592 | -35.2257 | -30.8273
20.5653 | -35.2261 | -30.8338
20.5714 | -35.2266 | -30.8404
20.5775 | -35.2271 | -30.8471
20.5836 | -35.2275 | -30.8536
20.5897 | -35.228 | -30.8605
20.5958 | -35.2285 | -30.867
20.6019 | -35.2289 | -30.8735
20.608 | -35.2293 | -30.8804
20.6141 | -35.2298 | -30.887
20.6202 | -35.2302 | -30.8938
20.6263 | -35.2306 | -30.9005
20.6324 | -35.2311 | -30.9072
20.6385 | -35.2315 | -30.914
20.6446 | -35.232 | -30.9208
20.6507 | -35.2324 | -30.9274
20.6568 | -35.2329 | -30.934
20.6629 | -35.2333 | -30.941
20.669 | -35.2338 | -30.9478
20.6751 | -35.2342 | -30.9545
20.6812 | -35.2346 | -30.9609
20.6873 | -35.235 | -30.9675
20.6934 | -35.2355 | -30.974
20.6995 | -35.236 | -30.9807
20.7057 | -35.2364 | -30.9873
20.7118 | -35.2369 | -30.9939
20.7179 | -35.2373 | -31.0003
20.724 | -35.2378 | -31.0072
20.7301 | -35.2382 | -31.014
20.7362 | -35.2387 | -31.0208
20.7423 | -35.2391 | -31.0274
20.7484 | -35.2396 | -31.0341
20.7545 | -35.24 | -31.0409
20.7606 | -35.2405 | -31.0477
20.7667 | -35.2409 | -31.0544
20.7728 | -35.2414 | -31.0615
20.7789 | -35.2419 | -31.0684
20.785 | -35.2424 | -31.075
20.7911 | -35.2428 | -31.082
20.7972 | -35.2433 | -31.0887
20.8033 | -35.2438 | -31.0955
20.8094 | -35.2442 | -31.1023
20.8155 | -35.2447 | -31.1087
20.8216 | -35.2452 | -31.1153
20.8277 | -35.2456 | -31.1221
20.8338 | -35.2461 | -31.1287
20.8399 | -35.2465 | -31.1351
20.8461 | -35.2469 | -31.1418
20.8522 | -35.2474 | -31.1489
20.8583 | -35.2478 | -31.1555
20.8644 | -35.2483 | -31.1623
20.8705 | -35.2487 | -31.1694
20.8766 | -35.2491 | -31.1763
20.8827 | -35.2495 | -31.183
20.8888 | -35.25 | -31.1894
20.8949 | -35.2504 | -31.1962
20.901 | -35.2508 | -31.2031
20.9071 | -35.2512 | -31.2096
20.9132 | -35.2516 | -31.2163
20.9193 | -35.2521 | -31.2231
20.9254 | -35.2525 | -31.2295
20.9315 | -35.253 | -31.2359
20.9376 | -35.2534 | -31.2428
20.9437 | -35.2538 | -31.2495
20.9498 | -35.2543 | -31.2558
20.9559 | -35.2547 | -31.2623
20.962 | -35.2551 | -31.2689
20.9681 | -35.2555 | -31.2753
20.9742 | -35.256 | -31.2822
20.9803 | -35.2565 | -31.2889
20.9864 | -35.2569 | -31.2956
20.9926 | -35.2573 | -31.3022
20.9987 | -35.2577 | -31.3088
21.0048 | -35.2582 | -31.3153
21.0109 | -35.2587 | -31.3221
21.017 | -35.2592 | -31.3289
21.0231 | -35.2596 | -31.3355
21.0292 | -35.2601 | -31.3424
21.0353 | -35.2605 | -31.349
21.0414 | -35.261 | -31.3556
21.0475 | -35.2614 | -31.3624
21.0536 | -35.2619 | -31.3689
21.0597 | -35.2623 | -31.3754
21.0658 | -35.2628 | -31.3819
21.0719 | -35.2632 | -31.3891
21.078 | -35.2637 | -31.3957
21.0841 | -35.2642 | -31.4028
21.0902 | -35.2648 | -31.4088
21.0963 | -35.2652 | -31.4164
21.1024 | -35.2657 | -31.4232
21.1085 | -35.2661 | -31.4297
21.1146 | -35.2666 | -31.4364
21.1207 | -35.2671 | -31.4435
21.1268 | -35.2676 | -31.4504
21.133 | -35.268 | -31.4571
21.1391 | -35.2684 | -31.4638
21.1452 | -35.2689 | -31.4705
21.1513 | -35.2693 | -31.4773
21.1574 | -35.2698 | -31.4837
21.1635 | -35.2703 | -31.4906
21.1696 | -35.2707 | -31.497
21.1757 | -35.2711 | -31.5035
21.1818 | -35.2716 | -31.5101
21.1879 | -35.272 | -31.5166
21.194 | -35.2724 | -31.5232
21.2001 | -35.2729 | -31.5296
21.2062 | -35.2733 | -31.5366
21.2123 | -35.2737 | -31.5425
21.2184 | -35.2741 | -31.5491
21.2245 | -35.2745 | -31.5555
21.2306 | -35.275 | -31.5623
21.2367 | -35.2754 | -31.5697
21.2428 | -35.2758 | -31.5761
21.2489 | -35.2762 | -31.583
21.255 | -35.2767 | -31.5893
21.2611 | -35.2771 | -31.5958
21.2672 | -35.2776 | -31.6024
21.2733 | -35.278 | -31.6088
21.2795 | -35.2785 | -31.6152
21.2856 | -35.2789 | -31.6217
21.2917 | -35.2794 | -31.6284
21.2978 | -35.2799 | -31.635
21.3039 | -35.2803 | -31.6417
21.31 | -35.2807 | -31.6485
21.3161 | -35.2811 | -31.6562
21.3222 | -35.2814 | -31.6632
21.3283 | -35.2819 | -31.6698
21.3344 | -35.2823 | -31.6764
21.3405 | -35.2827 | -31.6828
21.3466 | -35.2832 | -31.6899
21.3527 | -35.2836 | -31.6967
21.3588 | -35.2841 | -31.7034
21.3649 | -35.2846 | -31.7095
21.371 | -35.285 | -31.7161
21.3771 | -35.2855 | -31.7224
21.3832 | -35.2859 | -31.7296
21.3893 | -35.2864 | -31.7363
21.3954 | -35.2868 | -31.7432
21.4015 | -35.2873 | -31.7499
21.4076 | -35.2877 | -31.7568
21.4137 | -35.2882 | -31.7635
21.4199 | -35.2886 | -31.7705
21.426 | -35.2891 | -31.7773
21.4321 | -35.2895 | -31.7836
21.4382 | -35.29 | -31.7901
21.4443 | -35.2904 | -31.7968
21.4504 | -35.2909 | -31.8035
21.4565 | -35.2913 | -31.8106
21.4626 | -35.2918 | -31.8175
21.4687 | -35.2922 | -31.8244
21.4748 | -35.2926 | -31.8308
21.4809 | -35.2931 | -31.8374
21.487 | -35.2935 | -31.8446
21.4931 | -35.2939 | -31.8512
21.4992 | -35.2943 | -31.8579
21.5053 | -35.2948 | -31.8643
21.5114 | -35.2953 | -31.8707
21.5175 | -35.2957 | -31.8771
21.5236 | -35.2961 | -31.8838
21.5297 | -35.2965 | -31.8902
21.5358 | -35.2969 | -31.8967
21.5419 | -35.2973 | -31.9035
21.548 | -35.2977 | -31.9104
21.5541 | -35.2981 | -31.9171
21.5602 | -35.2985 | -31.9241
21.5664 | -35.2989 | -31.9308
21.5725 | -35.2994 | -31.9374
21.5786 | -35.2998 | -31.9439
21.5847 | -35.3002 | -31.9502
21.5908 | -35.3007 | -31.9567
21.5969 | -35.3012 | -31.9631
21.603 | -35.3016 | -31.9697
21.6091 | -35.3021 | -31.9767
21.6152 | -35.3025 | -31.9835
21.6213 | -35.303 | -31.9902
21.6274 | -35.3034 | -31.997
21.6335 | -35.3038 | -32.0034
21.6396 | -35.3042 | -32.0102
21.6457 | -35.3047 | -32.017
21.6518 | -35.3051 | -32.0235
21.6579 | -35.3055 | -32.03
21.664 | -35.3059 | -32.0367
21.6701 | -35.3063 | -32.0435
21.6762 | -35.3068 | -32.0501
21.6823 | -35.3072 | -32.057
21.6884 | -35.3076 | -32.0636
21.6945 | -35.308 | -32.0702
21.7006 | -35.3084 | -32.0772
21.7068 | -35.3088 | -32.0837
21.7129 | -35.3093 | -32.0901
21.719 | -35.3097 | -32.0968
21.7251 | -35.3101 | -32.1032
21.7312 | -35.3105 | -32.1101
21.7373 | -35.3109 | -32.1161
21.7434 | -35.3113 | -32.1229
21.7495 | -35.3118 | -32.1293
21.7556 | -35.3122 | -32.1357
21.7617 | -35.3126 | -32.1423
21.7678 | -35.313 | -32.1489
21.7739 | -35.3134 | -32.1553
21.78 | -35.3139 | -32.1618
21.7861 | -35.3143 | -32.1687
21.7922 | -35.3148 | -32.1753
21.7983 | -35.3152 | -32.1819
21.8044 | -35.3156 | -32.1887
21.8105 | -35.316 | -32.1954
21.8166 | -35.3165 | -32.2023
21.8227 | -35.3169 | -32.2088
21.8288 | -35.3173 | -32.2159
21.8349 | -35.3178 | -32.2227
21.841 | -35.3182 | -32.2296
21.8471 | -35.3187 | -32.236
21.8533 | -35.3191 | -32.2428
21.8594 | -35.3196 | -32.2493
21.8655 | -35.32 | -32.2558
21.8716 | -35.3205 | -32.2627
21.8777 | -35.3209 | -32.2695
21.8838 | -35.3213 | -32.2763
21.8899 | -35.3217 | -32.2827
21.896 | -35.3222 | -32.2891
21.9021 | -35.3226 | -32.2957
21.9082 | -35.323 | -32.3027
21.9143 | -35.3234 | -32.3097
21.9204 | -35.3239 | -32.316
21.9265 | -35.3243 | -32.3228
21.9326 | -35.3248 | -32.3295
21.9387 | -35.3252 | -32.3362
21.9448 | -35.3257 | -32.3429
21.9509 | -35.3262 | -32.3495
21.957 | -35.3266 | -32.3563
21.9631 | -35.3271 | -32.3628
21.9692 | -35.3275 | -32.3695
21.9753 | -35.3279 | -32.3761
21.9814 | -35.3283 | -32.3828
21.9875 | -35.3288 | -32.3891
21.9937 | -35.3292 | -32.3956
21.9998 | -35.3296 | -32.4021
22.0059 | -35.33 | -32.4087
22.012 | -35.3304 | -32.4153
22.0181 | -35.3309 | -32.422
22.0242 | -35.3313 | -32.4284
22.0303 | -35.3317 | -32.4349
22.0364 | -35.3321 | -32.4416
22.0425 | -35.3325 | -32.4482
22.0486 | -35.3329 | -32.4547
22.0547 | -35.3333 | -32.4614
22.0608 | -35.3337 | -32.4679
22.0669 | -35.3341 | -32.4744
22.073 | -35.3346 | -32.4805
22.0791 | -35.3351 | -32.4873
22.0852 | -35.3355 | -32.494
22.0913 | -35.3359 | -32.5006
22.0974 | -35.3364 | -32.5071
22.1035 | -35.3368 | -32.5136
22.1096 | -35.3372 | -32.5201
22.1157 | -35.3377 | -32.5265
22.1218 | -35.3381 | -32.5332
22.1279 | -35.3386 | -32.5395
22.134 | -35.339 | -32.546
22.1402 | -35.3394 | -32.5529
22.1463 | -35.3399 | -32.5595
22.1524 | -35.3402 | -32.5659
22.1585 | -35.3406 | -32.5723
22.1646 | -35.341 | -32.5786
22.1707 | -35.3414 | -32.5849
22.1768 | -35.3419 | -32.5915
22.1829 | -35.3424 | -32.5982
22.189 | -35.3428 | -32.6046
22.1951 | -35.3432 | -32.611
22.2012 | -35.3436 | -32.6174
22.2073 | -35.3441 | -32.6241
22.2134 | -35.3445 | -32.6309
22.2195 | -35.345 | -32.6375
22.2256 | -35.3454 | -32.6443
22.2317 | -35.3458 | -32.651
22.2378 | -35.3462 | -32.658
22.2439 | -35.3466 | -32.6645
22.25 | -35.3471 | -32.671
22.2561 | -35.3475 | -32.6773
22.2622 | -35.348 | -32.6838
22.2683 | -35.3484 | -32.6898
22.2744 | -35.3489 | -32.6961
22.2806 | -35.3493 | -32.7027
22.2867 | -35.3497 | -32.7091
22.2928 | -35.3502 | -32.7157
22.2989 | -35.3505 | -32.7221
22.305 | -35.3511 | -32.7289
22.3111 | -35.3515 | -32.7353
22.3172 | -35.3519 | -32.7418
22.3233 | -35.3523 | -32.7486
22.3294 | -35.3528 | -32.755
22.3355 | -35.3532 | -32.7615
22.3416 | -35.3536 | -32.7679
22.3477 | -35.354 | -32.7745
22.3538 | -35.3544 | -32.781
22.3599 | -35.3548 | -32.7878
22.366 | -35.3553 | -32.7942
22.3721 | -35.3558 | -32.8002
22.3782 | -35.3562 | -32.807
22.3843 | -35.3566 | -32.8138
22.3904 | -35.357 | -32.8206
22.3965 | -35.3574 | -32.8274
22.4026 | -35.3579 | -32.834
22.4087 | -35.3584 | -32.8401
22.4148 | -35.3588 | -32.8464
22.4209 | -35.3592 | -32.8529
22.4271 | -35.3597 | -32.8593
22.4332 | -35.3601 | -32.8658
22.4393 | -35.3605 | -32.8723
22.4454 | -35.3609 | -32.879
22.4515 | -35.3614 | -32.8854
22.4576 | -35.3619 | -32.8919
22.4637 | -35.3623 | -32.8985
22.4698 | -35.3627 | -32.9049
22.4759 | -35.3632 | -32.9112
22.482 | -35.3636 | -32.9176
22.4881 | -35.364 | -32.9241
22.4942 | -35.3644 | -32.9307
22.5003 | -35.3648 | -32.9373
22.5064 | -35.3653 | -32.9435
22.5125 | -35.3657 | -32.9498
22.5186 | -35.3661 | -32.9562
22.5247 | -35.3665 | -32.9624
22.5308 | -35.367 | -32.9692
22.5369 | -35.3674 | -32.9758
22.543 | -35.3678 | -32.982
22.5491 | -35.3683 | -32.9884
22.5552 | -35.3687 | -32.9948
22.5613 | -35.3691 | -33.0015
22.5675 | -35.3696 | -33.0079
22.5736 | -35.37 | -33.0144
22.5797 | -35.3704 | -33.0208
22.5858 | -35.3708 | -33.027
22.5919 | -35.3713 | -33.0335
22.598 | -35.3717 | -33.0401
22.6041 | -35.3721 | -33.0467
22.6102 | -35.3726 | -33.0531
22.6163 | -35.373 | -33.0594
22.6224 | -35.3733 | -33.0658
22.6285 | -35.3738 | -33.0725
22.6346 | -35.3742 | -33.0784
22.6407 | -35.3746 | -33.0852
22.6468 | -35.3751 | -33.0915
22.6529 | -35.3755 | -33.0969
22.659 | -35.3759 | -33.103
22.6651 | -35.3763 | -33.1097
22.6712 | -35.3767 | -33.1158
22.6773 | -35.3772 | -33.1224
22.6834 | -35.3776 | -33.1288
22.6895 | -35.3781 | -33.135
22.6956 | -35.3785 | -33.1414
22.7017 | -35.379 | -33.148
22.7079 | -35.3794 | -33.1544
22.714 | -35.3798 | -33.161
22.7201 | -35.3803 | -33.1669
22.7262 | -35.3807 | -33.1726
22.7323 | -35.3812 | -33.1787
22.7384 | -35.3816 | -33.1854
22.7445 | -35.382 | -33.1921
22.7506 | -35.3824 | -33.1986
22.7567 | -35.3828 | -33.2053
22.7628 | -35.3832 | -33.2115
22.7689 | -35.3836 | -33.2181
22.775 | -35.3841 | -33.2248
22.7811 | -35.3845 | -33.2315
22.7872 | -35.3849 | -33.2376
22.7933 | -35.3853 | -33.2438
22.7994 | -35.3857 | -33.25
22.8055 | -35.3862 | -33.2563
22.8116 | -35.3866 | -33.2626
22.8177 | -35.387 | -33.2694
22.8238 | -35.3874 | -33.2759
22.8299 | -35.3878 | -33.282
22.836 | -35.3882 | -33.288
22.8421 | -35.3886 | -33.2946
22.8482 | -35.3891 | -33.3011
22.8544 | -35.3895 | -33.3078
22.8605 | -35.39 | -33.3142
22.8666 | -35.3904 | -33.3204
22.8727 | -35.3908 | -33.3269
22.8788 | -35.3912 | -33.3334
22.8849 | -35.3917 | -33.34
22.891 | -35.3921 | -33.3461
22.8971 | -35.3925 | -33.3526
22.9032 | -35.393 | -33.3588
22.9093 | -35.3934 | -33.3649
22.9154 | -35.3939 | -33.3709
22.9215 | -35.3943 | -33.3773
22.9276 | -35.3947 | -33.3836
22.9337 | -35.395 | -33.3897
22.9398 | -35.3954 | -33.3958
22.9459 | -35.3959 | -33.4023
22.952 | -35.3963 | -33.4087
22.9581 | -35.3967 | -33.4151
22.9642 | -35.3971 | -33.4216
22.9703 | -35.3976 | -33.4278
22.9764 | -35.398 | -33.4342
22.9825 | -35.3984 | -33.4404
22.9886 | -35.3988 | -33.4469
22.9948 | -35.3992 | -33.4531
23.0009 | -35.3996 | -33.4594
23.007 | -35.4 | -33.465
23.0131 | -35.4004 | -33.4716
23.0192 | -35.4008 | -33.4776
23.0253 | -35.4013 | -33.484
23.0314 | -35.4017 | -33.4904
23.0375 | -35.4021 | -33.4967
23.0436 | -35.4026 | -33.5024
23.0497 | -35.403 | -33.509
23.0558 | -35.4035 | -33.5154
23.0619 | -35.4039 | -33.5218
23.068 | -35.4043 | -33.5283
23.0741 | -35.4047 | -33.5339
23.0802 | -35.4051 | -33.5403
23.0863 | -35.4055 | -33.5466
23.0924 | -35.406 | -33.5527
23.0985 | -35.4064 | -33.5588
23.1046 | -35.4069 | -33.5654
23.1107 | -35.4073 | -33.5713
23.1168 | -35.4077 | -33.5776
23.1229 | -35.4081 | -33.5835
23.129 | -35.4085 | -33.5897
23.1351 | -35.409 | -33.5961
23.1413 | -35.4094 | -33.6022
23.1474 | -35.4098 | -33.6088
23.1535 | -35.4102 | -33.6151
23.1596 | -35.4107 | -33.6211
23.1657 | -35.4112 | -33.6274
23.1718 | -35.4116 | -33.6335
23.1779 | -35.412 | -33.6398
23.184 | -35.4125 | -33.6455
23.1901 | -35.4129 | -33.6518
23.1962 | -35.4134 | -33.658
23.2023 | -35.4138 | -33.6644
23.2084 | -35.4143 | -33.6709
23.2145 | -35.4147 | -33.6773
23.2206 | -35.4151 | -33.6834
23.2267 | -35.4155 | -33.6896
23.2328 | -35.4159 | -33.6959
23.2389 | -35.4163 | -33.7021
23.245 | -35.4167 | -33.7083
23.2511 | -35.4171 | -33.7141
23.2572 | -35.4176 | -33.7204
23.2633 | -35.418 | -33.7264
23.2694 | -35.4184 | -33.733
23.2755 | -35.4189 | -33.7393
23.2817 | -35.4192 | -33.7452
23.2878 | -35.4197 | -33.7512
23.2939 | -35.4201 | -33.7574
23.3 | -35.4205 | -33.7633
23.3061 | -35.4209 | -33.7698
23.3122 | -35.4214 | -33.7757
23.3183 | -35.4219 | -33.7819
23.3244 | -35.4223 | -33.7878
23.3305 | -35.4227 | -33.7941
23.3366 | -35.4231 | -33.8006
23.3427 | -35.4236 | -33.8069
23.3488 | -35.424 | -33.8131
23.3549 | -35.4244 | -33.8194
23.361 | -35.4248 | -33.8251
23.3671 | -35.4252 | -33.8313
23.3732 | -35.4256 | -33.8373
23.3793 | -35.426 | -33.8433
23.3854 | -35.4265 | -33.8495
23.3915 | -35.4269 | -33.855
23.3976 | -35.4273 | -33.8612
23.4037 | -35.4278 | -33.8674
23.4098 | -35.4283 | -33.8738
23.4159 | -35.4287 | -33.8797
23.422 | -35.4291 | -33.886
23.4282 | -35.4294 | -33.8918
23.4343 | -35.4299 | -33.8979
23.4404 | -35.4303 | -33.9041
23.4465 | -35.4307 | -33.91
23.4526 | -35.4311 | -33.9164
23.4587 | -35.4315 | -33.9227
23.4648 | -35.4319 | -33.9287
23.4709 | -35.4323 | -33.9344
23.477 | -35.4327 | -33.9405
23.4831 | -35.4331 | -33.9464
23.4892 | -35.4335 | -33.9527
23.4953 | -35.4339 | -33.959
23.5014 | -35.4343 | -33.9651
23.5075 | -35.4347 | -33.9713
23.5136 | -35.4351 | -33.9778
23.5197 | -35.4355 | -33.984
23.5258 | -35.4359 | -33.9903
23.5319 | -35.4363 | -33.9963
23.538 | -35.4367 | -34.0023
23.5441 | -35.4372 | -34.0086
23.5502 | -35.4376 | -34.0148
23.5563 | -35.438 | -34.0211
23.5624 | -35.4384 | -34.027
23.5686 | -35.4388 | -34.0327
23.5747 | -35.4392 | -34.0387
23.5808 | -35.4396 | -34.0445
23.5869 | -35.44 | -34.0502
23.593 | -35.4405 | -34.0561
23.5991 | -35.4409 | -34.0628
23.6052 | -35.4414 | -34.0684
23.6113 | -35.4418 | -34.0749
23.6174 | -35.4422 | -34.0811
23.6235 | -35.4426 | -34.0873
23.6296 | -35.4431 | -34.0934
23.6357 | -35.4435 | -34.0995
23.6418 | -35.4439 | -34.1051
23.6479 | -35.4443 | -34.1111
23.654 | -35.4447 | -34.1169
23.6601 | -35.4451 | -34.1229
23.6662 | -35.4456 | -34.1287
23.6723 | -35.446 | -34.1345
23.6784 | -35.4464 | -34.1405
23.6845 | -35.4468 | -34.1467
23.6906 | -35.4473 | -34.1531
23.6967 | -35.4477 | -34.1591
23.7028 | -35.4481 | -34.1648
23.7089 | -35.4485 | -34.1708
23.7151 | -35.4489 | -34.1766
23.7212 | -35.4494 | -34.1829
23.7273 | -35.4498 | -34.1891
23.7334 | -35.4502 | -34.1952
23.7395 | -35.4506 | -34.2013
23.7456 | -35.451 | -34.2067
23.7517 | -35.4514 | -34.2121
23.7578 | -35.4518 | -34.218
23.7639 | -35.4522 | -34.2236
23.77 | -35.4527 | -34.2299
23.7761 | -35.4531 | -34.2358
23.7822 | -35.4536 | -34.2409
23.7883 | -35.4541 | -34.2472
23.7944 | -35.4545 | -34.2532
23.8005 | -35.4548 | -34.2595
23.8066 | -35.4553 | -34.2655
23.8127 | -35.4556 | -34.2707
23.8188 | -35.4561 | -34.2771
23.8249 | -35.4565 | -34.2829
23.831 | -35.4568 | -34.2888
23.8371 | -35.4572 | -34.295
23.8432 | -35.4576 | -34.301
23.8493 | -35.458 | -34.3063
23.8555 | -35.4584 | -34.3125
23.8616 | -35.4588 | -34.3185
23.8677 | -35.4592 | -34.3246
23.8738 | -35.4596 | -34.3305
23.8799 | -35.4601 | -34.3364
23.886 | -35.4605 | -34.3419
23.8921 | -35.4609 | -34.348
23.8982 | -35.4613 | -34.3537
23.9043 | -35.4616 | -34.3595
23.9104 | -35.462 | -34.3658
23.9165 | -35.4624 | -34.3715
23.9226 | -35.4628 | -34.3769
23.9287 | -35.4631 | -34.3824
23.9348 | -35.4635 | -34.3881
23.9409 | -35.4639 | -34.3941
23.947 | -35.4643 | -34.4
23.9531 | -35.4647 | -34.4052
23.9592 | -35.4651 | -34.4113
23.9653 | -35.4656 | -34.4172
23.9714 | -35.466 | -34.4235
23.9775 | -35.4665 | -34.4295
23.9836 | -35.4669 | -34.4356
23.9897 | -35.4673 | -34.441
23.9958 | -35.4678 | -34.4475
24.002 | -35.4682 | -34.454
24.0081 | -35.4686 | -34.4596
24.0142 | -35.469 | -34.4652
24.0203 | -35.4694 | -34.4713
24.0264 | -35.4698 | -34.4764
24.0325 | -35.4703 | -34.482
24.0386 | -35.4707 | -34.4882
24.0447 | -35.4711 | -34.4944
24.0508 | -35.4715 | -34.5005
24.0569 | -35.4719 | -34.5064
24.063 | -35.4723 | -34.5122
24.0691 | -35.4726 | -34.5175
24.0752 | -35.473 | -34.5233
24.0813 | -35.4735 | -34.5292
24.0874 | -35.4738 | -34.5356
24.0935 | -35.4743 | -34.5417
24.0996 | -35.4747 | -34.5471
24.1057 | -35.4751 | -34.5528
24.1118 | -35.4755 | -34.5589
24.1179 | -35.4759 | -34.5648
24.124 | -35.4763 | -34.5704
24.1301 | -35.4767 | -34.576
24.1362 | -35.4771 | -34.5821
24.1424 | -35.4776 | -34.5883
24.1485 | -35.478 | -34.5945
24.1546 | -35.4785 | -34.6009
24.1607 | -35.4788 | -34.6068
24.1668 | -35.4792 | -34.6122
24.1729 | -35.4797 | -34.6182
24.179 | -35.4801 | -34.6239
24.1851 | -35.4804 | -34.6299
24.1912 | -35.4808 | -34.6359
24.1973 | -35.4812 | -34.6419
24.2034 | -35.4815 | -34.6472
24.2095 | -35.482 | -34.6528
24.2156 | -35.4825 | -34.6589
24.2217 | -35.4828 | -34.6647
24.2278 | -35.4833 | -34.6705
24.2339 | -35.4836 | -34.6763
24.24 | -35.4841 | -34.6813
24.2461 | -35.4846 | -34.6871
24.2522 | -35.4849 | -34.6931
24.2583 | -35.4854 | -34.6988
24.2644 | -35.4857 | -34.7048
24.2705 | -35.4862 | -34.7104
24.2766 | -35.4866 | -34.7162
24.2827 | -35.487 | -34.7218
24.2889 | -35.4873 | -34.7278
24.295 | -35.4877 | -34.7335
24.3011 | -35.4881 | -34.7394
24.3072 | -35.4885 | -34.7447
24.3133 | -35.4889 | -34.7506
24.3194 | -35.4893 | -34.7566
24.3255 | -35.4897 | -34.7622
24.3316 | -35.4901 | -34.7681
24.3377 | -35.4905 | -34.7744
24.3438 | -35.4909 | -34.7798
24.3499 | -35.4913 | -34.7857
24.356 | -35.4917 | -34.791
24.3621 | -35.4921 | -34.7967
24.3682 | -35.4925 | -34.8026
24.3743 | -35.4929 | -34.8085
24.3804 | -35.4933 | -34.8139
24.3865 | -35.4937 | -34.8195
24.3926 | -35.4942 | -34.8256
24.3987 | -35.4946 | -34.8312
24.4048 | -35.495 | -34.8373
24.4109 | -35.4955 | -34.843
24.417 | -35.4959 | -34.8485
24.4231 | -35.4963 | -34.8539
24.4293 | -35.4967 | -34.8594
24.4354 | -35.4971 | -34.8649
24.4415 | -35.4975 | -34.8706
24.4476 | -35.498 | -34.8755
24.4537 | -35.4985 | -34.8811
24.4598 | -35.499 | -34.8865
24.4659 | -35.4994 | -34.8924
24.472 | -35.4998 | -34.8982
24.4781 | -35.5002 | -34.9035
24.4842 | -35.5006 | -34.9085
24.4903 | -35.5011 | -34.9145
24.4964 | -35.5015 | -34.9199
24.5025 | -35.5019 | -34.9254
24.5086 | -35.5023 | -34.9307
24.5147 | -35.5027 | -34.9363
24.5208 | -35.5031 | -34.941
24.5269 | -35.5035 | -34.9464
24.533 | -35.5039 | -34.952
24.5391 | -35.5044 | -34.9573
24.5452 | -35.5048 | -34.9627
24.5513 | -35.5053 | -34.9682
24.5574 | -35.5057 | -34.974
24.5635 | -35.5061 | -34.9798
24.5696 | -35.5065 | -34.9853
24.5758 | -35.5069 | -34.9907
24.5819 | -35.5074 | -34.9963
24.588 | -35.5078 | -35.0014
24.5941 | -35.5082 | -35.0066
24.6002 | -35.5087 | -35.0128
24.6063 | -35.5092 | -35.0177
24.6124 | -35.5096 | -35.0229
24.6185 | -35.51 | -35.0287
24.6246 | -35.5104 | -35.0344
24.6307 | -35.5109 | -35.0403
24.6368 | -35.5114 | -35.0456
24.6429 | -35.5118 | -35.0518
24.649 | -35.5122 | -35.0572
24.6551 | -35.5126 | -35.0629
24.6612 | -35.513 | -35.0682
24.6673 | -35.5134 | -35.0736
24.6734 | -35.5139 | -35.079
24.6795 | -35.5143 | -35.0846
24.6856 | -35.5147 | -35.0904
24.6917 | -35.5151 | -35.0955
24.6978 | -35.5156 | -35.1007
24.7039 | -35.5159 | -35.1062
24.71 | -35.5163 | -35.1113
24.7162 | -35.5167 | -35.1167
24.7223 | -35.5171 | -35.1224
24.7284 | -35.5175 | -35.128
24.7345 | -35.5179 | -35.1336
24.7406 | -35.5183 | -35.139
24.7467 | -35.5187 | -35.1445
24.7528 | -35.5191 | -35.1498
24.7589 | -35.5196 | -35.1547
24.765 | -35.52 | -35.1602
24.7711 | -35.5205 | -35.1653
24.7772 | -35.5209 | -35.1703
24.7833 | -35.5213 | -35.1762
24.7894 | -35.5217 | -35.181
24.7955 | -35.5222 | -35.1864
24.8016 | -35.5226 | -35.1923
24.8077 | -35.5231 | -35.1979
24.8138 | -35.5235 | -35.2037
24.8199 | -35.524 | -35.2095
24.826 | -35.5243 | -35.2148
24.8321 | -35.5247 | -35.2202
24.8382 | -35.525 | -35.2257
24.8443 | -35.5253 | -35.2317
24.8504 | -35.5257 | -35.2375
24.8565 | -35.5261 | -35.2423
24.8627 | -35.5265 | -35.2482
24.8688 | -35.5268 | -35.2544
24.8749 | -35.5272 | -35.2595
24.881 | -35.5276 | -35.2652
24.8871 | -35.5281 | -35.2708
24.8932 | -35.5285 | -35.2763
24.8993 | -35.5289 | -35.2815
24.9054 | -35.5293 | -35.2862
24.9115 | -35.5297 | -35.2919
24.9176 | -35.5302 | -35.2973
24.9237 | -35.5306 | -35.3025
24.9298 | -35.531 | -35.3078
24.9359 | -35.5314 | -35.3133
24.942 | -35.5318 | -35.3186
24.9481 | -35.5323 | -35.3239
24.9542 | -35.5327 | -35.3292
24.9603 | -35.5331 | -35.3346
24.9664 | -35.5335 | -35.3402
24.9725 | -35.5339 | -35.3456
24.9786 | -35.5343 | -35.3515
24.9847 | -35.5347 | -35.3568
24.9908 | -35.5351 | -35.3622
24.9969 | -35.5355 | -35.3672
25.0031 | -35.5359 | -35.3725
25.0092 | -35.5363 | -35.3777
25.0153 | -35.5367 | -35.383
25.0214 | -35.5371 | -35.3889
25.0275 | -35.5376 | -35.3937
25.0336 | -35.538 | -35.3988
25.0397 | -35.5384 | -35.404
25.0458 | -35.5389 | -35.4093
25.0519 | -35.5393 | -35.4153
25.058 | -35.5397 | -35.4208
25.0641 | -35.5401 | -35.426
25.0702 | -35.5404 | -35.431
25.0763 | -35.5408 | -35.4367
25.0824 | -35.5413 | -35.441
25.0885 | -35.5417 | -35.4463
25.0946 | -35.5421 | -35.452
25.1007 | -35.5425 | -35.4575
25.1068 | -35.5429 | -35.4628
25.1129 | -35.5434 | -35.4685
25.119 | -35.5438 | -35.4729
25.1251 | -35.5442 | -35.4782
25.1312 | -35.5446 | -35.4836
25.1373 | -35.545 | -35.4893
25.1435 | -35.5454 | -35.4944
25.1496 | -35.5458 | -35.4992
25.1557 | -35.5462 | -35.5048
25.1618 | -35.5466 | -35.5101
25.1679 | -35.547 | -35.5156
25.174 | -35.5474 | -35.5211
25.1801 | -35.5478 | -35.5262
25.1862 | -35.5482 | -35.5313
25.1923 | -35.5486 | -35.5366
25.1984 | -35.549 | -35.542
25.2045 | -35.5494 | -35.5477
25.2106 | -35.5498 | -35.5533
25.2167 | -35.5502 | -35.5587
25.2228 | -35.5505 | -35.5642
25.2289 | -35.5509 | -35.5696
25.235 | -35.5514 | -35.5749
25.2411 | -35.5517 | -35.5802
25.2472 | -35.5521 | -35.5853
25.2533 | -35.5525 | -35.5908
25.2594 | -35.5529 | -35.5957
25.2655 | -35.5533 | -35.6013
25.2716 | -35.5537 | -35.6072
25.2777 | -35.5541 | -35.6124
25.2838 | -35.5546 | -35.6176
25.29 | -35.555 | -35.6233
25.2961 | -35.5553 | -35.6284
25.3022 | -35.5556 | -35.6337
25.3083 | -35.5559 | -35.639
25.3144 | -35.5563 | -35.6445
25.3205 | -35.5567 | -35.6499
25.3266 | -35.5572 | -35.6548
25.3327 | -35.5575 | -35.6597
25.3388 | -35.5579 | -35.6652
25.3449 | -35.5583 | -35.6708
25.351 | -35.5586 | -35.6769
25.3571 | -35.5591 | -35.6822
25.3632 | -35.5595 | -35.6874
25.3693 | -35.5599 | -35.6928
25.3754 | -35.5602 | -35.698
25.3815 | -35.5606 | -35.7032
25.3876 | -35.561 | -35.7085
25.3937 | -35.5614 | -35.7137
25.3998 | -35.5618 | -35.7191
25.4059 | -35.5622 | -35.7249
25.412 | -35.5625 | -35.7304
25.4181 | -35.5629 | -35.7356
25.4242 | -35.5633 | -35.7409
25.4304 | -35.5638 | -35.7465
25.4365 | -35.5642 | -35.7516
25.4426 | -35.5646 | -35.7569
25.4487 | -35.5649 | -35.7617
25.4548 | -35.5653 | -35.7666
25.4609 | -35.5658 | -35.7718
25.467 | -35.5662 | -35.7766
25.4731 | -35.5665 | -35.7816
25.4792 | -35.5669 | -35.7871
25.4853 | -35.5673 | -35.7924
25.4914 | -35.5677 | -35.7978
25.4975 | -35.5681 | -35.8029
25.5036 | -35.5685 | -35.8082
25.5097 | -35.5689 | -35.8136
25.5158 | -35.5692 | -35.819
25.5219 | -35.5696 | -35.8245
25.528 | -35.57 | -35.8298
25.5341 | -35.5705 | -35.8351
25.5402 | -35.5708 | -35.8401
25.5463 | -35.5711 | -35.8457
25.5524 | -35.5715 | -35.8511
25.5585 | -35.5719 | -35.8566
25.5646 | -35.5723 | -35.862
25.5707 | -35.5728 | -35.8672
25.5769 | -35.5731 | -35.8721
25.583 | -35.5735 | -35.877
25.5891 | -35.5739 | -35.8819
25.5952 | -35.5743 | -35.8874
25.6013 | -35.5747 | -35.8928
25.6074 | -35.575 | -35.8978
25.6135 | -35.5754 | -35.9032
25.6196 | -35.5758 | -35.9084
25.6257 | -35.5762 | -35.9134
25.6318 | -35.5766 | -35.9186
25.6379 | -35.5769 | -35.9238
25.644 | -35.5773 | -35.9286
25.6501 | -35.5777 | -35.9338
25.6562 | -35.5781 | -35.9388
25.6623 | -35.5785 | -35.9441
25.6684 | -35.5789 | -35.9496
25.6745 | -35.5793 | -35.955
25.6806 | -35.5798 | -35.9596
25.6867 | -35.5802 | -35.9648
25.6928 | -35.5806 | -35.97
25.6989 | -35.581 | -35.9751
25.705 | -35.5814 | -35.9805
25.7111 | -35.5818 | -35.9857
25.7173 | -35.5821 | -35.9904
25.7234 | -35.5825 | -35.9955
25.7295 | -35.5829 | -36.0005
25.7356 | -35.5833 | -36.0053
25.7417 | -35.5838 | -36.0105
25.7478 | -35.5842 | -36.0154
25.7539 | -35.5845 | -36.0204
25.76 | -35.5849 | -36.0257
25.7661 | -35.5853 | -36.0309
25.7722 | -35.5857 | -36.0362
25.7783 | -35.5861 | -36.0415
25.7844 | -35.5865 | -36.0468
25.7905 | -35.5868 | -36.0519
25.7966 | -35.5872 | -36.0572
25.8027 | -35.5876 | -36.0624
25.8088 | -35.588 | -36.0677
25.8149 | -35.5884 | -36.0729
25.821 | -35.5887 | -36.0781
25.8271 | -35.5891 | -36.0832
25.8332 | -35.5895 | -36.0884
25.8393 | -35.5899 | -36.0936
25.8454 | -35.5903 | -36.0989
25.8515 | -35.5907 | -36.1042
25.8576 | -35.591 | -36.109
25.8638 | -35.5914 | -36.1142
25.8699 | -35.5917 | -36.1195
25.876 | -35.5922 | -36.1244
25.8821 | -35.5925 | -36.1293
25.8882 | -35.5929 | -36.134
25.8943 | -35.5933 | -36.1391
25.9004 | -35.5937 | -36.144
25.9065 | -35.5941 | -36.1491
25.9126 | -35.5944 | -36.1542
25.9187 | -35.5948 | -36.1593
25.9248 | -35.5952 | -36.1643
25.9309 | -35.5956 | -36.1694
25.937 | -35.5959 | -36.1748
25.9431 | -35.5963 | -36.1801
25.9492 | -35.5967 | -36.1852
25.9553 | -35.597 | -36.1903
25.9614 | -35.5974 | -36.1948
25.9675 | -35.5978 | -36.1998
25.9736 | -35.5982 | -36.205
25.9797 | -35.5986 | -36.2103
25.9858 | -35.599 | -36.2153
25.9919 | -35.5993 | -36.2208
25.998 | -35.5997 | -36.2253
26.0042 | -35.6001 | -36.2304
26.0103 | -35.6005 | -36.2355
26.0164 | -35.6008 | -36.2406
26.0225 | -35.6012 | -36.2457
26.0286 | -35.6016 | -36.2506
26.0347 | -35.602 | -36.2555
26.0408 | -35.6024 | -36.2607
26.0469 | -35.6028 | -36.2658
26.053 | -35.6031 | -36.2712
26.0591 | -35.6036 | -36.2761
26.0652 | -35.6038 | -36.2809
26.0713 | -35.6042 | -36.2859
26.0774 | -35.6046 | -36.2909
26.0835 | -35.605 | -36.2963
26.0896 | -35.6054 | -36.3014
26.0957 | -35.6058 | -36.3065
26.1018 | -35.6062 | -36.3112
26.1079 | -35.6066 | -36.3163
26.114 | -35.607 | -36.3211
26.1201 | -35.6073 | -36.3258
26.1262 | -35.6077 | -36.3312
26.1323 | -35.608 | -36.3362
26.1384 | -35.6084 | -36.3408
26.1445 | -35.6088 | -36.3457
26.1507 | -35.6092 | -36.3507
26.1568 | -35.6096 | -36.3557
26.1629 | -35.61 | -36.3605
26.169 | -35.6103 | -36.3655
26.1751 | -35.6107 | -36.3703
26.1812 | -35.6111 | -36.3754
26.1873 | -35.6115 | -36.3806
26.1934 | -35.6118 | -36.3859
26.1995 | -35.6122 | -36.3912
26.2056 | -35.6125 | -36.3959
26.2117 | -35.6129 | -36.4011
26.2178 | -35.6132 | -36.406
26.2239 | -35.6136 | -36.4108
26.23 | -35.614 | -36.416
26.2361 | -35.6144 | -36.4212
26.2422 | -35.6147 | -36.4259
26.2483 | -35.6151 | -36.4309
26.2544 | -35.6155 | -36.4358
26.2605 | -35.6158 | -36.441
26.2666 | -35.6162 | -36.4459
26.2727 | -35.6166 | -36.451
26.2788 | -35.6169 | -36.4563
26.2849 | -35.6173 | -36.4615
26.2911 | -35.6177 | -36.4664
26.2972 | -35.6181 | -36.4715
26.3033 | -35.6185 | -36.4765
26.3094 | -35.6189 | -36.4816
26.3155 | -35.6192 | -36.4865
26.3216 | -35.6196 | -36.4916
26.3277 | -35.6199 | -36.4967
26.3338 | -35.6203 | -36.5019
26.3399 | -35.6207 | -36.5069
26.346 | -35.621 | -36.5118
26.3521 | -35.6214 | -36.517
26.3582 | -35.6218 | -36.5218
26.3643 | -35.6221 | -36.5269
26.3704 | -35.6225 | -36.532
26.3765 | -35.6228 | -36.537
26.3826 | -35.6232 | -36.5415
26.3887 | -35.6236 | -36.5467
26.3948 | -35.624 | -36.5515
26.4009 | -35.6244 | -36.5565
26.407 | -35.6248 | -36.5614
26.4131 | -35.6252 | -36.5664
26.4192 | -35.6256 | -36.5711
26.4253 | -35.626 | -36.5761
26.4314 | -35.6263 | -36.5811
26.4376 | -35.6267 | -36.5861
26.4437 | -35.627 | -36.591
26.4498 | -35.6274 | -36.5962
26.4559 | -35.6277 | -36.6009
26.462 | -35.628 | -36.6059
26.4681 | -35.6284 | -36.6108
26.4742 | -35.6288 | -36.6159
26.4803 | -35.6292 | -36.6209
26.4864 | -35.6295 | -36.6259
26.4925 | -35.6299 | -36.6309
26.4986 | -35.6303 | -36.6361
26.5047 | -35.6307 | -36.6411
26.5108 | -35.631 | -36.6463
26.5169 | -35.6313 | -36.6513
26.523 | -35.6317 | -36.6559
26.5291 | -35.6321 | -36.6608
26.5352 | -35.6324 | -36.666
26.5413 | -35.6328 | -36.671
26.5474 | -35.6331 | -36.6761
26.5535 | -35.6335 | -36.6812
26.5596 | -35.6338 | -36.6863
26.5657 | -35.6342 | -36.6914
26.5718 | -35.6346 | -36.6964
26.578 | -35.6349 | -36.7014
26.5841 | -35.6353 | -36.7065
26.5902 | -35.6356 | -36.7118
26.5963 | -35.6359 | -36.7166
26.6024 | -35.6363 | -36.7215
26.6085 | -35.6366 | -36.7267
26.6146 | -35.637 | -36.7317
26.6207 | -35.6374 | -36.7366
26.6268 | -35.6377 | -36.7415
26.6329 | -35.638 | -36.7465
26.639 | -35.6383 | -36.7517
26.6451 | -35.6387 | -36.7565
26.6512 | -35.639 | -36.7615
26.6573 | -35.6394 | -36.7665
26.6634 | -35.6397 | -36.7711
26.6695 | -35.64 | -36.776
26.6756 | -35.6404 | -36.781
26.6817 | -35.6407 | -36.7859
26.6878 | -35.6411 | -36.7909
26.6939 | -35.6414 | -36.7959
26.7 | -35.6418 | -36.8007
26.7061 | -35.6421 | -36.806
26.7122 | -35.6424 | -36.8109
26.7183 | -35.6428 | -36.8157
26.7245 | -35.6431 | -36.8207
26.7306 | -35.6435 | -36.8256
26.7367 | -35.6438 | -36.8306
26.7428 | -35.6442 | -36.8352
26.7489 | -35.6446 | -36.8402
26.755 | -35.6449 | -36.8454
26.7611 | -35.6453 | -36.8505
26.7672 | -35.6456 | -36.8552
26.7733 | -35.646 | -36.8601
26.7794 | -35.6464 | -36.865
26.7855 | -35.6468 | -36.8698
26.7916 | -35.6471 | -36.8746
26.7977 | -35.6475 | -36.8795
26.8038 | -35.6478 | -36.8842
26.8099 | -35.6481 | -36.889
26.816 | -35.6485 | -36.8939
26.8221 | -35.6488 | -36.8987
26.8282 | -35.6491 | -36.9039
26.8343 | -35.6494 | -36.9089
26.8404 | -35.6498 | -36.9137
26.8465 | -35.6501 | -36.9188
26.8526 | -35.6505 | -36.9236
26.8587 | -35.6508 | -36.9284
26.8649 | -35.6512 | -36.9333
26.871 | -35.6515 | -36.9381
26.8771 | -35.6519 | -36.9429
26.8832 | -35.6522 | -36.9478
26.8893 | -35.6525 | -36.9527
26.8954 | -35.6529 | -36.9578
26.9015 | -35.6533 | -36.9627
26.9076 | -35.6536 | -36.9676
26.9137 | -35.6539 | -36.9722
26.9198 | -35.6542 | -36.977
26.9259 | -35.6545 | -36.9821
26.932 | -35.6549 | -36.9869
26.9381 | -35.6553 | -36.9917
26.9442 | -35.6556 | -36.9966
26.9503 | -35.656 | -37.0018
26.9564 | -35.6563 | -37.0065
26.9625 | -35.6567 | -37.0114
26.9686 | -35.6571 | -37.0162
26.9747 | -35.6574 | -37.0209
26.9808 | -35.6578 | -37.0254
26.9869 | -35.6581 | -37.0304
26.993 | -35.6585 | -37.0354
26.9991 | -35.6588 | -37.0404
27.0052 | -35.6592 | -37.0457
27.0114 | -35.6595 | -37.0505
27.0175 | -35.6598 | -37.0549
27.0236 | -35.6601 | -37.0599
27.0297 | -35.6605 | -37.0649
27.0358 | -35.6608 | -37.07
27.0419 | -35.6612 | -37.0749
27.048 | -35.6615 | -37.0797
27.0541 | -35.6619 | -37.0847
27.0602 | -35.6622 | -37.0896
27.0663 | -35.6625 | -37.0944
27.0724 | -35.6628 | -37.0994
27.0785 | -35.6632 | -37.1044
27.0846 | -35.6635 | -37.109
27.0907 | -35.6639 | -37.1142
27.0968 | -35.6642 | -37.1189
27.1029 | -35.6646 | -37.1237
27.109 | -35.6649 | -37.1285
27.1151 | -35.6652 | -37.1335
27.1212 | -35.6655 | -37.1384
27.1273 | -35.6658 | -37.1433
27.1334 | -35.6662 | -37.1484
27.1395 | -35.6665 | -37.1533
27.1456 | -35.6669 | -37.1581
27.1518 | -35.6672 | -37.1631
27.1579 | -35.6676 | -37.1676
27.164 | -35.6679 | -37.1726
27.1701 | -35.6682 | -37.1776
27.1762 | -35.6685 | -37.1824
27.1823 | -35.6688 | -37.1873
27.1884 | -35.6692 | -37.1923
27.1945 | -35.6695 | -37.1972
27.2006 | -35.6698 | -37.2021
27.2067 | -35.6701 | -37.2071
27.2128 | -35.6705 | -37.2119
27.2189 | -35.6708 | -37.2166
27.225 | -35.6711 | -37.2217
27.2311 | -35.6714 | -37.2261
27.2372 | -35.6718 | -37.2311
27.2433 | -35.6721 | -37.2358
27.2494 | -35.6724 | -37.2408
27.2555 | -35.6728 | -37.2457
27.2616 | -35.6731 | -37.2506
27.2677 | -35.6734 | -37.2554
27.2738 | -35.6737 | -37.2603
27.2799 | -35.6741 | -37.2652
27.286 | -35.6745 | -37.2702
27.2921 | -35.6748 | -37.275
27.2983 | -35.6751 | -37.2797
27.3044 | -35.6755 | -37.2845
27.3105 | -35.6759 | -37.2895
27.3166 | -35.6762 | -37.2944
27.3227 | -35.6766 | -37.2991
27.3288 | -35.6769 | -37.304
27.3349 | -35.6772 | -37.3085
27.341 | -35.6775 | -37.3135
27.3471 | -35.6778 | -37.3182
27.3532 | -35.6781 | -37.3229
27.3593 | -35.6785 | -37.3278
27.3654 | -35.6788 | -37.3327
27.3715 | -35.6791 | -37.3374
27.3776 | -35.6795 | -37.3422
27.3837 | -35.6799 | -37.3471
27.3898 | -35.6802 | -37.3518
27.3959 | -35.6805 | -37.3565
27.402 | -35.6808 | -37.3612
27.4081 | -35.6812 | -37.3663
27.4142 | -35.6815 | -37.3714
27.4203 | -35.6819 | -37.3766
27.4264 | -35.6821 | -37.3811
27.4325 | -35.6825 | -37.3859
27.4387 | -35.6828 | -37.3909
27.4448 | -35.6831 | -37.3956
27.4509 | -35.6835 | -37.4004
27.457 | -35.6838 | -37.4054
27.4631 | -35.6841 | -37.4101
27.4692 | -35.6844 | -37.4148
27.4753 | -35.6847 | -37.4195
27.4814 | -35.685 | -37.4244
27.4875 | -35.6854 | -37.4292
27.4936 | -35.6857 | -37.4338
27.4997 | -35.6861 | -37.4385
27.5058 | -35.6864 | -37.4433
27.5119 | -35.6867 | -37.448
27.518 | -35.687 | -37.4527
27.5241 | -35.6873 | -37.4575
27.5302 | -35.6876 | -37.4621
27.5363 | -35.688 | -37.4669
27.5424 | -35.6883 | -37.4712
27.5485 | -35.6886 | -37.4761
27.5546 | -35.6889 | -37.4808
27.5607 | -35.6892 | -37.4857
27.5668 | -35.6895 | -37.4901
27.5729 | -35.6899 | -37.4949
27.5791 | -35.6902 | -37.4998
27.5852 | -35.6905 | -37.5046
27.5913 | -35.6908 | -37.5092
27.5974 | -35.6911 | -37.514
27.6035 | -35.6915 | -37.5191
27.6096 | -35.6918 | -37.5239
27.6157 | -35.6921 | -37.5285
27.6218 | -35.6924 | -37.5333
27.6279 | -35.6927 | -37.538
27.634 | -35.693 | -37.5431
27.6401 | -35.6933 | -37.5481
27.6462 | -35.6937 | -37.553
27.6523 | -35.6939 | -37.5578
27.6584 | -35.6943 | -37.5627
27.6645 | -35.6946 | -37.5676
27.6706 | -35.695 | -37.5724
27.6767 | -35.6953 | -37.5772
27.6828 | -35.6956 | -37.5818
27.6889 | -35.6959 | -37.5869
27.695 | -35.6962 | -37.5918
27.7011 | -35.6965 | -37.5965
27.7072 | -35.6968 | -37.6014
27.7133 | -35.6971 | -37.6062
27.7194 | -35.6973 | -37.6109
27.7256 | -35.6976 | -37.6159
27.7317 | -35.6979 | -37.6208
27.7378 | -35.6982 | -37.6257
27.7439 | -35.6986 | -37.6304
27.75 | -35.6989 | -37.6355
27.7561 | -35.6992 | -37.6399
27.7622 | -35.6995 | -37.6447
27.7683 | -35.6998 | -37.6492
27.7744 | -35.7002 | -37.654
27.7805 | -35.7005 | -37.6587
27.7866 | -35.7008 | -37.6634
27.7927 | -35.701 | -37.6678
27.7988 | -35.7014 | -37.6725
27.8049 | -35.7017 | -37.6775
27.811 | -35.7021 | -37.6824
27.8171 | -35.7025 | -37.6874
27.8232 | -35.7028 | -37.6921
27.8293 | -35.703 | -37.6968
27.8354 | -35.7034 | -37.7016
27.8415 | -35.7037 | -37.7063
27.8476 | -35.7041 | -37.7113
27.8537 | -35.7044 | -37.7161
27.8598 | -35.7047 | -37.7208
27.866 | -35.705 | -37.7255
27.8721 | -35.7053 | -37.7302
27.8782 | -35.7056 | -37.7351
27.8843 | -35.7059 | -37.7399
27.8904 | -35.7063 | -37.7448
27.8965 | -35.7066 | -37.7494
27.9026 | -35.7069 | -37.7545
27.9087 | -35.7072 | -37.7595
27.9148 | -35.7075 | -37.764
27.9209 | -35.7078 | -37.7688
27.927 | -35.7081 | -37.7733
27.9331 | -35.7084 | -37.7782
27.9392 | -35.7087 | -37.783
27.9453 | -35.709 | -37.788
27.9514 | -35.7093 | -37.7927
27.9575 | -35.7096 | -37.7975
27.9636 | -35.71 | -37.8022
27.9697 | -35.7102 | -37.8072
27.9758 | -35.7105 | -37.8123
27.9819 | -35.7108 | -37.817
27.988 | -35.7111 | -37.8216
27.9941 | -35.7114 | -37.8266
28.0002 | -35.7117 | -37.8312
28.0063 | -35.7121 | -37.8364
28.0125 | -35.7123 | -37.841
28.0186 | -35.7126 | -37.8457
28.0247 | -35.7129 | -37.8506
28.0308 | -35.7132 | -37.8555
28.0369 | -35.7135 | -37.8603
28.043 | -35.7138 | -37.8651
28.0491 | -35.7142 | -37.8699
28.0552 | -35.7145 | -37.8748
28.0613 | -35.7148 | -37.8796
28.0674 | -35.7151 | -37.8844
28.0735 | -35.7154 | -37.8891

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4721 | -42.3391
21.4908 | -35.4825 | -34.9084
16.4948 | -34.8652 | -27.5905
12.6602 | -34.502 | -21.3503
9.7171 | -34.2871 | -16.4634
7.4582 | -34.1411 | -12.5002
5.7244 | -34.0351 | -9.2893
4.3936 | -33.9694 | -6.6941
3.3722 | -33.951 | -4.6489
2.5883 | -33.9451 | -3.0978
1.9866 | -33.9458 | -1.8622
1.5248 | -33.9555 | -0.86629
1.1703 | -33.958 | -0.22246
1 | -33.958 | 0.10136
0.89822 | -33.9578 | 0.37263
0.68941 | -33.9391 | 0.95216
0.52919 | -33.9452 | 1.3971
0.40612 | -33.9414 | 2.0055
0.31171 | -33.9415 | 2.4532
0.23928 | -33.9334 | 3.1796
0.18358 | -33.9392 | 3.5958
0.14091 | -33.9676 | 4.4682
0.10819 | -33.9857 | 5.6268
0.083017 | -34.006 | 7.1056
0.063705 | -34.0581 | 8.9783
0.048923 | -34.121 | 11.4767
0.037479 | -34.2256 | 14.6654
0.028801 | -34.4073 | 18.6953
0.022078 | -34.715 | 24.0905
0.017023 | -35.1998 | 30.5519
0.013018 | -36.0134 | 38.2521
0.010014 | -37.1314 | 46.1618

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003743 | -11.8513
0.007485 | -35.2363
0.011228 | -40.0166
0.014971 | -43.4592
0.018713 | -45.2164
0.022456 | -47.0824
0.026199 | -48.6433
0.029942 | -50.0952
0.033684 | -51.4765
0.037427 | -52.9741
0.04117 | -53.9824
0.044912 | -55.2116
0.048655 | -55.7604
0.052398 | -55.8451
0.05614 | -56.4352
0.059883 | -57.3511
0.063626 | -58.5941
0.067368 | -59.7139
0.071111 | -60.4334
0.074854 | -61.205
0.078596 | -61.7938
0.082339 | -62.484
0.086082 | -63.4022
0.089825 | -64.0009
0.093567 | -64.1986
0.09731 | -64.2125
0.10105 | -64.1735
0.10479 | -64.8782
0.10854 | -65.738
0.11228 | -66.1141
0.11602 | -66.1913
0.11977 | -65.735
0.12351 | -65.0851
0.12725 | -66.2088
0.13099 | -66.9786
0.13474 | -68.121
0.13848 | -68.7095
0.14222 | -69.3026
0.14597 | -68.4553
0.14971 | -67.8491
0.15345 | -68.4814
0.15719 | -69.3325
0.16094 | -69.4582
0.16468 | -70.2605
0.16842 | -70.3051
0.17216 | -70.4846
0.17591 | -69.5959
0.17965 | -69.671
0.18339 | -70.0425
0.18713 | -70.4918
0.19088 | -70.8172
0.19462 | -70.8606
0.19836 | -71.0067
0.20211 | -72.1256
0.20585 | -72.1504
0.20959 | -72.1948
0.21333 | -72.6684
0.21708 | -72.3709
0.22082 | -73.0205
0.22456 | -73.2519
0.2283 | -72.9819
0.23205 | -72.8047
0.23579 | -72.5935
0.23953 | -72.6221
0.24327 | -73.3142
0.24702 | -74.0018
0.25076 | -75.1149
0.2545 | -75.1481
0.25825 | -75.0204
0.26199 | -74.8983
0.26573 | -73.9731
0.26947 | -75.7906
0.27322 | -75.3155
0.27696 | -75.2925
0.2807 | -75.6252
0.28444 | -75.4666
0.28819 | -75.8284
0.29193 | -75.515
0.29567 | -75.1702
0.29941 | -75.5986
0.30316 | -75.758
0.3069 | -75.2941
0.31064 | -75.6021
0.31439 | -75.8972
0.31813 | -76.6275
0.32187 | -76.9132
0.32561 | -76.6997
0.32936 | -77.0158
0.3331 | -76.6907
0.33684 | -76.5677
0.34059 | -76.626
0.34433 | -76.2337
0.34807 | -75.9765
0.35181 | -76.5044
0.35556 | -78.4231
0.3593 | -78.9418
0.36304 | -77.7552
0.36678 | -77.1997
0.37053 | -77.8178
0.37427 | -76.9906
0.37801 | -77.0543
0.38175 | -79.1963
0.3855 | -78.409
0.38924 | -77.9163
0.39298 | -78.4928
0.39672 | -78.2813
0.40047 | -78.8471
0.40421 | -79.1778
0.40795 | -79.8562
0.4117 | -79.2698
0.41544 | -79.2166
0.41918 | -79.274
0.42292 | -78.569
0.42667 | -79.3003
0.43041 | -78.991
0.43415 | -80.1596
0.43789 | -80.1484
0.44164 | -80.6776
0.44538 | -80.8708
0.44912 | -80.805
0.45287 | -79.8143
0.45661 | -80.0972
0.46035 | -80.2778
0.46409 | -80.0788
0.46784 | -79.8357
0.47158 | -79.7829
0.47532 | -80.0349
0.47906 | -80.4509
0.48281 | -80.5669
0.48655 | -80.493
0.49029 | -79.8325
0.49404 | -80.2882
0.49778 | -80.6015
0.50152 | -80.6013
0.50526 | -80.6678
0.50901 | -80.4386
0.51275 | -81.5484
0.51649 | -80.7347
0.52023 | -79.8776
0.52398 | -81.7005
0.52772 | -82.1651
0.53146 | -81.4601
0.53521 | -81.4533
0.53895 | -81.8704
0.54269 | -82.0462
0.54643 | -81.6914
0.55017 | -81.639
0.55392 | -81.896
0.55766 | -82.3363
0.5614 | -82.7949
0.56515 | -82.54
0.56889 | -82.2167
0.57263 | -82.1762
0.57637 | -82.0796
0.58012 | -81.6964
0.58386 | -81.637
0.5876 | -82.0959
0.59135 | -82.475
0.59509 | -82.9295
0.59883 | -83.6702
0.60257 | -83.5619
0.60632 | -82.9134
0.61006 | -82.555
0.6138 | -82.5402
0.61754 | -82.9116
0.62129 | -83.5671
0.62503 | -83.6272
0.62877 | -83.0865
0.63252 | -82.6358
0.63626 | -82.3134
0.64 | -82.3712
0.64374 | -82.7697
0.64748 | -82.893
0.65123 | -82.9201
0.65497 | -83.0515
0.65871 | -83.4345
0.66246 | -83.6231
0.6662 | -83.8109
0.66994 | -84.2646
0.67368 | -84.7702
0.67743 | -84.6197
0.68117 | -83.6938
0.68491 | -83.1005
0.68866 | -83.167
0.6924 | -83.4971
0.69614 | -83.8022
0.69988 | -83.825
0.70363 | -83.6946
0.70737 | -84.1439
0.71111 | -84.7782
0.71485 | -85.2221
0.7186 | -85.6091
0.72234 | -85.4782
0.72608 | -84.7988
0.72982 | -84.6278
0.73357 | -84.5549
0.73731 | -84.0022
0.74105 | -83.8906
0.74479 | -84.5318
0.74854 | -85.0571
0.75228 | -85.115
0.75602 | -85.2223
0.75977 | -84.9657
0.76351 | -84.8066
0.76725 | -85.1513
0.77099 | -85.3967
0.77474 | -85.4415
0.77848 | -85.3099
0.78222 | -85.2379
0.78597 | -85.1569
0.78971 | -85.0023
0.79345 | -84.8468
0.79719 | -84.7293
0.80094 | -84.4335
0.80468 | -84.2502
0.80842 | -84.7214
0.81216 | -85.5614
0.81591 | -85.9792
0.81965 | -86.0138
0.82339 | -85.9747
0.82713 | -86.0464
0.83088 | -86.0885
0.83462 | -85.7049
0.83836 | -85.541
0.8421 | -85.8591
0.84585 | -85.7949
0.84959 | -85.492
0.85333 | -85.6154
0.85708 | -85.902
0.86082 | -86.2745
0.86456 | -86.8273
0.8683 | -86.8948
0.87205 | -86.6658
0.87579 | -86.5279
0.87953 | -86.5957
0.88328 | -86.7564
0.88702 | -86.8084
0.89076 | -86.7775
0.8945 | -86.6561
0.89825 | -86.45
0.90199 | -86.5094
0.90573 | -86.6782
0.90947 | -86.96
0.91322 | -87.1259
0.91696 | -87.074
0.9207 | -87.0496
0.92444 | -87.1301
0.92819 | -87.3286
0.93193 | -87.4698
0.93567 | -87.2402
0.93942 | -87.0505
0.94316 | -87.2378
0.9469 | -87.5917
0.95064 | -87.7661
0.95439 | -87.4287
0.95813 | -87.211
0.96187 | -87.5239
0.96561 | -87.7916
0.96936 | -87.5718
0.9731 | -87.2899
0.97684 | -86.9426
0.98059 | -86.8024
0.98433 | -87.1366
0.98807 | -87.2373
0.99181 | -86.7081
0.99556 | -86.5699
0.9993 | -86.9314
1.003 | -87.2104
1.0068 | -87.2295
1.0105 | -87.1675
1.0143 | -87.3931
1.018 | -87.9079
1.0218 | -88.0506
1.0255 | -87.6775
1.0292 | -87.4132
1.033 | -87.3289
1.0367 | -87.4054
1.0405 | -87.7744
1.0442 | -87.6183
1.048 | -87.3809
1.0517 | -88.0517
1.0554 | -88.4102
1.0592 | -87.6503
1.0629 | -87.0738
1.0667 | -87.1718
1.0704 | -87.8936
1.0742 | -88.7393
1.0779 | -88.7063
1.0816 | -88.6242
1.0854 | -88.4876
1.0891 | -88.3638
1.0929 | -88.2017
1.0966 | -88.007
1.1004 | -87.8722
1.1041 | -87.8522
1.1078 | -87.91
1.1116 | -88.0783
1.1153 | -88.4381
1.1191 | -88.8882
1.1228 | -89.1009
1.1265 | -88.9861
1.1303 | -88.9371
1.134 | -89.1908
1.1378 | -89.5233
1.1415 | -89.606
1.1453 | -89.4659
1.149 | -89.3985
1.1527 | -89.4428
1.1565 | -89.2794
1.1602 | -88.9075
1.164 | -88.6681
1.1677 | -88.7125
1.1715 | -88.8749
1.1752 | -88.8937
1.1789 | -88.8406
1.1827 | -88.9286
1.1864 | -89.1565
1.1902 | -89.3937
1.1939 | -89.4633
1.1977 | -89.3128
1.2014 | -89.0861
1.2051 | -88.9577
1.2089 | -89.0387
1.2126 | -89.2424
1.2164 | -89.3262
1.2201 | -89.221
1.2239 | -89.0899
1.2276 | -88.9744
1.2313 | -88.921
1.2351 | -89.0153
1.2388 | -89.1314
1.2426 | -89.1469
1.2463 | -89.1248
1.2501 | -89.0677
1.2538 | -89.0137
1.2575 | -89.0782
1.2613 | -89.257
1.265 | -89.366
1.2688 | -89.3068
1.2725 | -89.2208
1.2763 | -89.2593
1.28 | -89.3499
1.2837 | -89.3725
1.2875 | -89.374
1.2912 | -89.4881
1.295 | -89.7156
1.2987 | -89.8447
1.3025 | -89.7712
1.3062 | -89.7424
1.3099 | -89.9128
1.3137 | -90.1076
1.3174 | -90.0998
1.3212 | -89.8411
1.3249 | -89.4794
1.3287 | -89.2971
1.3324 | -89.4129
1.3361 | -89.62
1.3399 | -89.6925
1.3436 | -89.7187
1.3474 | -89.9327
1.3511 | -90.3544
1.3549 | -90.6566
1.3586 | -90.6124
1.3623 | -90.5107
1.3661 | -90.6327
1.3698 | -90.9097
1.3736 | -91.1258
1.3773 | -91.0658
1.3811 | -90.7685
1.3848 | -90.5234
1.3885 | -90.5184
1.3923 | -90.764
1.396 | -91.1225
1.3998 | -91.3498
1.4035 | -91.3033
1.4073 | -90.9194
1.411 | -90.2594
1.4147 | -89.657
1.4185 | -89.4515
1.4222 | -89.685
1.426 | -90.2012
1.4297 | -90.7937
1.4334 | -91.1318
1.4372 | -91.0737
1.4409 | -90.9761
1.4447 | -91.076
1.4484 | -91.205
1.4522 | -91.1717
1.4559 | -91.0312
1.4596 | -90.965
1.4634 | -90.9394
1.4671 | -90.7573
1.4709 | -90.3927
1.4746 | -90.0696
1.4784 | -90.0013
1.4821 | -90.1537
1.4858 | -90.3792
1.4896 | -90.6652
1.4933 | -90.9868
1.4971 | -91.2947
1.5008 | -91.5851
1.5046 | -91.6634
1.5083 | -91.4668
1.512 | -91.2869
1.5158 | -91.1566
1.5195 | -90.9295
1.5233 | -90.6288
1.527 | -90.409
1.5308 | -90.4048
1.5345 | -90.5928
1.5382 | -90.8877
1.542 | -91.172
1.5457 | -91.2651
1.5495 | -91.1846
1.5532 | -91.108
1.557 | -91.105
1.5607 | -91.143
1.5644 | -91.207
1.5682 | -91.3169
1.5719 | -91.44
1.5757 | -91.5366
1.5794 | -91.6817
1.5832 | -91.8897
1.5869 | -92.0287
1.5906 | -92.0209
1.5944 | -91.8772
1.5981 | -91.6538
1.6019 | -91.509
1.6056 | -91.6024
1.6094 | -91.8335
1.6131 | -92.0366
1.6168 | -92.0338
1.6206 | -91.9904
1.6243 | -91.9067
1.6281 | -91.9117
1.6318 | -92.06
1.6356 | -92.2192
1.6393 | -92.2104
1.643 | -92.0044
1.6468 | -91.7256
1.6505 | -91.5109
1.6543 | -91.4361
1.658 | -91.51
1.6618 | -91.6615
1.6655 | -91.7956
1.6692 | -91.8662
1.673 | -91.8691
1.6767 | -91.8372
1.6805 | -91.8478
1.6842 | -91.9057
1.688 | -91.9478
1.6917 | -91.9519
1.6954 | -91.9933
1.6992 | -92.1413
1.7029 | -92.3271
1.7067 | -92.3713
1.7104 | -92.1978
1.7142 | -91.9585
1.7179 | -91.853
1.7216 | -91.9369
1.7254 | -92.1264
1.7291 | -92.3376
1.7329 | -92.5232
1.7366 | -92.6666
1.7404 | -92.72
1.7441 | -92.6716
1.7478 | -92.6201
1.7516 | -92.6833
1.7553 | -92.8839
1.7591 | -93.1182
1.7628 | -93.2317
1.7666 | -93.1765
1.7703 | -92.9669
1.774 | -92.7033
1.7778 | -92.48
1.7815 | -92.3511
1.7853 | -92.3283
1.789 | -92.3899
1.7927 | -92.5038
1.7965 | -92.6255
1.8002 | -92.6902
1.804 | -92.6473
1.8077 | -92.4999
1.8115 | -92.338
1.8152 | -92.2367
1.8189 | -92.1957
1.8227 | -92.1918
1.8264 | -92.2361
1.8302 | -92.3704
1.8339 | -92.6155
1.8377 | -92.9039
1.8414 | -93.1024
1.8451 | -93.129
1.8489 | -93.0653
1.8526 | -93.0503
1.8564 | -93.1384
1.8601 | -93.2854
1.8639 | -93.3779
1.8676 | -93.3724
1.8713 | -93.342
1.8751 | -93.3713
1.8788 | -93.4149
1.8826 | -93.3924
1.8863 | -93.3499
1.8901 | -93.3661
1.8938 | -93.4026
1.8975 | -93.3456
1.9013 | -93.1252
1.905 | -92.8589
1.9088 | -92.7128
1.9125 | -92.7395
1.9163 | -92.8597
1.92 | -92.9878
1.9237 | -93.132
1.9275 | -93.2847
1.9312 | -93.3893
1.935 | -93.4025
1.9387 | -93.3346
1.9425 | -93.2678
1.9462 | -93.2721
1.9499 | -93.3374
1.9537 | -93.4207
1.9574 | -93.5029
1.9612 | -93.6128
1.9649 | -93.7509
1.9687 | -93.8595
1.9724 | -93.8614
1.9761 | -93.7944
1.9799 | -93.7909
1.9836 | -93.9105
1.9874 | -94.064
1.9911 | -94.1397
1.9949 | -94.1067
1.9986 | -93.9505
2.0023 | -93.6646
2.0061 | -93.2965
2.0098 | -92.9664
2.0136 | -92.8204
2.0173 | -92.9079
2.0211 | -93.1281
2.0248 | -93.3246
2.0285 | -93.4155
2.0323 | -93.4402
2.036 | -93.4433
2.0398 | -93.4347
2.0435 | -93.4534
2.0473 | -93.5325
2.051 | -93.6019
2.0547 | -93.5699
2.0585 | -93.5021
2.0622 | -93.5399
2.066 | -93.6901
2.0697 | -93.8773
2.0734 | -94.0655
2.0772 | -94.2267
2.0809 | -94.2518
2.0847 | -94.0898
2.0884 | -93.8732
2.0922 | -93.7409
2.0959 | -93.7161
2.0996 | -93.7782
2.1034 | -93.9561
2.1071 | -94.2418
2.1109 | -94.5115
2.1146 | -94.6652
2.1184 | -94.7127
2.1221 | -94.7325
2.1258 | -94.7858
2.1296 | -94.8025
2.1333 | -94.6841
2.1371 | -94.4276
2.1408 | -94.1401
2.1446 | -93.9364
2.1483 | -93.8686
2.152 | -93.9287
2.1558 | -94.0133
2.1595 | -93.8868
2.1633 | -93.7233
2.167 | -93.4942
2.1708 | -93.2693
2.1745 | -93.1401
2.1782 | -93.1682
2.182 | -93.3599
2.1857 | -93.6622
2.1895 | -93.9671
2.1932 | -94.1732
2.197 | -94.2388
2.2007 | -94.2094
2.2044 | -94.1794
2.2082 | -94.1823
2.2119 | -94.2163
2.2157 | -94.2341
2.2194 | -94.2171
2.2232 | -94.1644
2.2269 | -94.088
2.2306 | -94.0029
2.2344 | -93.9553
2.2381 | -93.9723
2.2419 | -94.0536
2.2456 | -94.1612
2.2494 | -94.265
2.2531 | -94.3639
2.2568 | -94.4856
2.2606 | -94.6161
2.2643 | -94.7541
2.2681 | -94.8768
2.2718 | -94.9587
2.2756 | -94.9509
2.2793 | -94.8248
2.283 | -94.5994
2.2868 | -94.3484
2.2905 | -94.1481
2.2943 | -94.0393
2.298 | -94.0261
2.3018 | -94.0897
2.3055 | -94.1662
2.3092 | -94.2062
2.313 | -94.1774
2.3167 | -94.1113
2.3205 | -94.0706
2.3242 | -94.0957
2.328 | -94.2077
2.3317 | -94.407
2.3354 | -94.6914
2.3392 | -95.0087
2.3429 | -95.2636
2.3467 | -95.3826
2.3504 | -95.3619
2.3542 | -95.261
2.3579 | -95.1304
2.3616 | -95.0033
2.3654 | -94.8923
2.3691 | -94.7961
2.3729 | -94.6951
2.3766 | -94.5838
2.3804 | -94.5104
2.3841 | -94.5207
2.3878 | -94.6304
2.3916 | -94.7981
2.3953 | -94.9618
2.3991 | -95.0587
2.4028 | -95.0757
2.4066 | -95.0236
2.4103 | -94.9374
2.414 | -94.8479
2.4178 | -94.7905
2.4215 | -94.7738
2.4253 | -94.8131
2.429 | -94.8806
2.4327 | -94.9375
2.4365 | -94.9438
2.4402 | -94.8956
2.444 | -94.8162
2.4477 | -94.7697
2.4515 | -94.7924
2.4552 | -94.8701
2.4589 | -94.9409
2.4627 | -94.9938
2.4664 | -95.0232
2.4702 | -95.0528
2.4739 | -95.0875
2.4777 | -95.1264
2.4814 | -95.1751
2.4851 | -95.1982
2.4889 | -95.1419
2.4926 | -94.9734
2.4964 | -94.7325
2.5001 | -94.5114
2.5039 | -94.3772
2.5076 | -94.3619
2.5113 | -94.4421
2.5151 | -94.5608
2.5188 | -94.6808
2.5226 | -94.7914
2.5263 | -94.9064
2.5301 | -95.0204
2.5338 | -95.1179
2.5375 | -95.1405
2.5413 | -95.0591
2.545 | -94.9203
2.5488 | -94.7881
2.5525 | -94.7205
2.5563 | -94.7143
2.56 | -94.7186
2.5637 | -94.7014
2.5675 | -94.6739
2.5712 | -94.6783
2.575 | -94.7538
2.5787 | -94.9389
2.5825 | -95.2252
2.5862 | -95.5783
2.5899 | -95.9529
2.5937 | -96.2513
2.5974 | -96.4129
2.6012 | -96.4255
2.6049 | -96.3171
2.6087 | -96.1551
2.6124 | -95.9831
2.6161 | -95.8159
2.6199 | -95.6589
2.6236 | -95.5301
2.6274 | -95.4592
2.6311 | -95.4436
2.6349 | -95.4508
2.6386 | -95.4208
2.6423 | -95.3122
2.6461 | -95.1609
2.6498 | -95.0312
2.6536 | -94.9912
2.6573 | -95.0575
2.6611 | -95.1721
2.6648 | -95.2517
2.6685 | -95.2623
2.6723 | -95.2546
2.676 | -95.3009
2.6798 | -95.4269
2.6835 | -95.5845
2.6873 | -95.6925
2.691 | -95.7117
2.6947 | -95.675
2.6985 | -95.6931
2.7022 | -95.7307
2.706 | -95.7921
2.7097 | -95.8446
2.7134 | -95.8541
2.7172 | -95.8021
2.7209 | -95.6837
2.7247 | -95.5444
2.7284 | -95.4402
2.7322 | -95.4068
2.7359 | -95.4425
2.7396 | -95.5035
2.7434 | -95.5585
2.7471 | -95.5867
2.7509 | -95.5975
2.7546 | -95.6331
2.7584 | -95.7299
2.7621 | -95.8866
2.7658 | -96.0837
2.7696 | -96.2673
2.7733 | -96.3687
2.7771 | -96.3504
2.7808 | -96.2286
2.7846 | -96.0645
2.7883 | -95.9128
2.792 | -95.8123
2.7958 | -95.7942
2.7995 | -95.8365
2.8033 | -95.8942
2.807 | -95.9334
2.8108 | -95.9483
2.8145 | -95.9504
2.8182 | -95.9639
2.822 | -95.9957
2.8257 | -96.0506
2.8295 | -96.1175
2.8332 | -96.1871
2.837 | -96.2484
2.8407 | -96.2844
2.8444 | -96.286
2.8482 | -96.2661
2.8519 | -96.2578
2.8557 | -96.2769
2.8594 | -96.3066
2.8632 | -96.3279
2.8669 | -96.3176
2.8706 | -96.2675
2.8744 | -96.1947
2.8781 | -96.1156
2.8819 | -96.045
2.8856 | -95.9964
2.8894 | -95.9728
2.8931 | -95.9934
2.8968 | -96.0802
2.9006 | -96.1916
2.9043 | -96.2868
2.9081 | -96.3483
2.9118 | -96.3428
2.9156 | -96.2907
2.9193 | -96.225
2.923 | -96.1691
2.9268 | -96.1394
2.9305 | -96.1397
2.9343 | -96.1679
2.938 | -96.2187
2.9418 | -96.2729
2.9455 | -96.3027
2.9492 | -96.2992
2.953 | -96.2498
2.9567 | -96.164
2.9605 | -96.0847
2.9642 | -96.0391
2.968 | -96.0573
2.9717 | -96.139
2.9754 | -96.2402
2.9792 | -96.2967
2.9829 | -96.2687
2.9867 | -96.1715
2.9904 | -96.0415
2.9942 | -95.9287
2.9979 | -95.8533
3.0016 | -95.8242
3.0054 | -95.825
3.0091 | -95.8436
3.0129 | -95.8481
3.0166 | -95.8241
3.0204 | -95.7805
3.0241 | -95.7488
3.0278 | -95.7633
3.0316 | -95.8494
3.0353 | -95.9963
3.0391 | -96.1943
3.0428 | -96.3966
3.0465 | -96.5674
3.0503 | -96.6653
3.054 | -96.6822
3.0578 | -96.6395
3.0615 | -96.573
3.0653 | -96.5145
3.069 | -96.4919
3.0727 | -96.5147
3.0765 | -96.5693
3.0802 | -96.6305
3.084 | -96.6728
3.0877 | -96.6838
3.0915 | -96.675
3.0952 | -96.6519
3.0989 | -96.641
3.1027 | -96.6563
3.1064 | -96.6981
3.1102 | -96.7852
3.1139 | -96.9076
3.1177 | -97.0594
3.1214 | -97.2142
3.1251 | -97.3393
3.1289 | -97.3897
3.1326 | -97.3478
3.1364 | -97.242
3.1401 | -97.1163
3.1439 | -97.0146
3.1476 | -96.952
3.1513 | -96.936
3.1551 | -96.9701
3.1588 | -97.0436
3.1626 | -97.1631
3.1663 | -97.3141
3.1701 | -97.4577
3.1738 | -97.5446
3.1775 | -97.5406
3.1813 | -97.4599
3.185 | -97.3225
3.1888 | -97.1885
3.1925 | -97.0983
3.1963 | -97.0638
3.2 | -97.0788
3.2037 | -97.13
3.2075 | -97.1974
3.2112 | -97.2372
3.215 | -97.2348
3.2187 | -97.2091
3.2225 | -97.1931
3.2262 | -97.2267
3.2299 | -97.3118
3.2337 | -97.4579
3.2374 | -97.6203
3.2412 | -97.7537
3.2449 | -97.8589
3.2487 | -97.9204
3.2524 | -97.9282
3.2561 | -97.8773
3.2599 | -97.7989
3.2636 | -97.6961
3.2674 | -97.5716
3.2711 | -97.4291
3.2749 | -97.2876
3.2786 | -97.1472
3.2823 | -97.032
3.2861 | -96.939
3.2898 | -96.8692
3.2936 | -96.829
3.2973 | -96.797
3.3011 | -96.7794
3.3048 | -96.7797
3.3085 | -96.8048
3.3123 | -96.8542
3.316 | -96.9203
3.3198 | -97.0148
3.3235 | -97.1362
3.3273 | -97.2775
3.331 | -97.4274
3.3347 | -97.5818
3.3385 | -97.7174
3.3422 | -97.8191
3.346 | -97.8784
3.3497 | -97.8865
3.3535 | -97.85
3.3572 | -97.7738
3.3609 | -97.669
3.3647 | -97.5636
3.3684 | -97.473
3.3722 | -97.3979
3.3759 | -97.3585
3.3796 | -97.3315
3.3834 | -97.3059
3.3871 | -97.2822
3.3909 | -97.2732
3.3946 | -97.2771
3.3984 | -97.2887
3.4021 | -97.3208
3.4058 | -97.3712
3.4096 | -97.433
3.4133 | -97.4867
3.4171 | -97.5336
3.4208 | -97.5618
3.4246 | -97.5887
3.4283 | -97.6306
3.432 | -97.6809
3.4358 | -97.7573
3.4395 | -97.8537
3.4433 | -97.9458
3.447 | -98.0199
3.4508 | -98.0983
3.4545 | -98.189
3.4582 | -98.2669
3.462 | -98.337
3.4657 | -98.3806
3.4695 | -98.3789
3.4732 | -98.31
3.477 | -98.1785
3.4807 | -98.0031
3.4844 | -97.8277
3.4882 | -97.6776
3.4919 | -97.5768
3.4957 | -97.5232
3.4994 | -97.5191
3.5032 | -97.5546
3.5069 | -97.5982
3.5106 | -97.6465
3.5144 | -97.6877
3.5181 | -97.7206
3.5219 | -97.7405
3.5256 | -97.7467
3.5294 | -97.7482
3.5331 | -97.7474
3.5368 | -97.7492
3.5406 | -97.7467
3.5443 | -97.7524
3.5481 | -97.7628
3.5518 | -97.7858
3.5556 | -97.7958
3.5593 | -97.7849
3.563 | -97.7566
3.5668 | -97.7036
3.5705 | -97.6416
3.5743 | -97.5838
3.578 | -97.5543
3.5818 | -97.5684
3.5855 | -97.6217
3.5892 | -97.7016
3.593 | -97.792
3.5967 | -97.882
3.6005 | -97.9531
3.6042 | -97.9948
3.608 | -98.0121
3.6117 | -98.0053
3.6154 | -97.9777
3.6192 | -97.9449
3.6229 | -97.9317
3.6267 | -97.9533
3.6304 | -98.0055
3.6342 | -98.0775
3.6379 | -98.1525
3.6416 | -98.2214
3.6454 | -98.2772
3.6491 | -98.3209
3.6529 | -98.3554
3.6566 | -98.3966
3.6604 | -98.4261
3.6641 | -98.4369
3.6678 | -98.4366
3.6716 | -98.4157
3.6753 | -98.3618
3.6791 | -98.2862
3.6828 | -98.2079
3.6865 | -98.1316
3.6903 | -98.0674
3.694 | -98.0183
3.6978 | -97.9905
3.7015 | -97.9884
3.7053 | -98.0021
3.709 | -98.0367
3.7127 | -98.0938
3.7165 | -98.1747
3.7202 | -98.2563
3.724 | -98.3219
3.7277 | -98.3672
3.7315 | -98.4005
3.7352 | -98.4378
3.7389 | -98.4949
3.7427 | -98.5681
3.7464 | -98.639
3.7502 | -98.6939
3.7539 | -98.7234
3.7577 | -98.7246
3.7614 | -98.6907
3.7651 | -98.6368
3.7689 | -98.5672
3.7726 | -98.496
3.7764 | -98.427
3.7801 | -98.356
3.7839 | -98.3012
3.7876 | -98.2707
3.7913 | -98.266
3.7951 | -98.2751
3.7988 | -98.2901
3.8026 | -98.3142
3.8063 | -98.3266
3.8101 | -98.3205
3.8138 | -98.3105
3.8175 | -98.3014
3.8213 | -98.312
3.825 | -98.3271
3.8288 | -98.3443
3.8325 | -98.3623
3.8363 | -98.3929
3.84 | -98.4338
3.8437 | -98.4725
3.8475 | -98.5017
3.8512 | -98.5264
3.855 | -98.5486
3.8587 | -98.5617
3.8625 | -98.5755
3.8662 | -98.5949
3.8699 | -98.6304
3.8737 | -98.691
3.8774 | -98.7575
3.8812 | -98.8374
3.8849 | -98.9116
3.8887 | -98.9687
3.8924 | -98.998
3.8961 | -98.9978
3.8999 | -98.9699
3.9036 | -98.9211
3.9074 | -98.8631
3.9111 | -98.7967
3.9149 | -98.7343
3.9186 | -98.6925
3.9223 | -98.6856
3.9261 | -98.7069
3.9298 | -98.7501
3.9336 | -98.8001
3.9373 | -98.8487
3.9411 | -98.8824
3.9448 | -98.8882
3.9485 | -98.8767
3.9523 | -98.8481
3.956 | -98.8105
3.9598 | -98.765
3.9635 | -98.6985
3.9673 | -98.6284
3.971 | -98.567
3.9747 | -98.5185
3.9785 | -98.4964
3.9822 | -98.5057
3.986 | -98.5348
3.9897 | -98.5877
3.9935 | -98.6521
3.9972 | -98.727
4.0009 | -98.8026
4.0047 | -98.8887
4.0084 | -98.9783
4.0122 | -99.0627
4.0159 | -99.1534
4.0196 | -99.2432
4.0234 | -99.3252
4.0271 | -99.4054
4.0309 | -99.4624
4.0346 | -99.4764
4.0384 | -99.4463
4.0421 | -99.3484
4.0458 | -99.2164
4.0496 | -99.0877
4.0533 | -98.9683
4.0571 | -98.8716
4.0608 | -98.798
4.0646 | -98.7601
4.0683 | -98.7557
4.072 | -98.7826
4.0758 | -98.8329
4.0795 | -98.8886
4.0833 | -98.9412
4.087 | -98.9697
4.0908 | -98.957
4.0945 | -98.9181
4.0982 | -98.8685
4.102 | -98.8268
4.1057 | -98.8073
4.1095 | -98.8072
4.1132 | -98.8284
4.117 | -98.876
4.1207 | -98.9472
4.1244 | -99.0254
4.1282 | -99.0867
4.1319 | -99.1385
4.1357 | -99.174
4.1394 | -99.1913
4.1432 | -99.1966
4.1469 | -99.1925
4.1506 | -99.1931
4.1544 | -99.2084
4.1581 | -99.2283
4.1619 | -99.2599
4.1656 | -99.3051
4.1694 | -99.3635
4.1731 | -99.4151
4.1768 | -99.4627
4.1806 | -99.503
4.1843 | -99.539
4.1881 | -99.5801
4.1918 | -99.6326
4.1956 | -99.6846
4.1993 | -99.7497
4.203 | -99.82
4.2068 | -99.8707
4.2105 | -99.895
4.2143 | -99.8953
4.218 | -99.8659
4.2218 | -99.8144
4.2255 | -99.7508
4.2292 | -99.6816
4.233 | -99.6187
4.2367 | -99.5632
4.2405 | -99.5143
4.2442 | -99.479
4.248 | -99.4576
4.2517 | -99.4451
4.2554 | -99.4384
4.2592 | -99.4358
4.2629 | -99.4174
4.2667 | -99.3925
4.2704 | -99.3752
4.2742 | -99.354
4.2779 | -99.3406
4.2816 | -99.3467
4.2854 | -99.3709
4.2891 | -99.4179
4.2929 | -99.4808
4.2966 | -99.5496
4.3004 | -99.6327
4.3041 | -99.7208
4.3078 | -99.8048
4.3116 | -99.8846
4.3153 | -99.9549
4.3191 | -99.9926
4.3228 | -100.0307
4.3266 | -100.0392
4.3303 | -100.0283
4.334 | -99.9919
4.3378 | -99.9148
4.3415 | -99.8188
4.3453 | -99.7001
4.349 | -99.5707
4.3527 | -99.4563
4.3565 | -99.3621
4.3602 | -99.2967
4.364 | -99.2819
4.3677 | -99.3076
4.3715 | -99.3594
4.3752 | -99.4262
4.3789 | -99.4973
4.3827 | -99.5667
4.3864 | -99.6224
4.3902 | -99.6679
4.3939 | -99.6956
4.3977 | -99.7253
4.4014 | -99.7425
4.4051 | -99.7645
4.4089 | -99.8044
4.4126 | -99.8462
4.4164 | -99.8983
4.4201 | -99.9523
4.4239 | -99.9938
4.4276 | -100.032
4.4313 | -100.0599
4.4351 | -100.0483
4.4388 | -100.0052
4.4426 | -99.9414
4.4463 | -99.8588
4.4501 | -99.7502
4.4538 | -99.625
4.4575 | -99.5089
4.4613 | -99.4005
4.465 | -99.3232
4.4688 | -99.2681
4.4725 | -99.2446
4.4763 | -99.2373
4.48 | -99.243
4.4837 | -99.2618
4.4875 | -99.29
4.4912 | -99.3327
4.495 | -99.3843
4.4987 | -99.4448
4.5025 | -99.5148
4.5062 | -99.585
4.5099 | -99.6496
4.5137 | -99.693
4.5174 | -99.7156
4.5212 | -99.725
4.5249 | -99.714
4.5287 | -99.6859
4.5324 | -99.6564
4.5361 | -99.6205
4.5399 | -99.5921
4.5436 | -99.5705
4.5474 | -99.5613
4.5511 | -99.5695
4.5549 | -99.5912
4.5586 | -99.616
4.5623 | -99.6357
4.5661 | -99.647
4.5698 | -99.637
4.5736 | -99.6162
4.5773 | -99.6033
4.5811 | -99.5984
4.5848 | -99.6241
4.5885 | -99.6896
4.5923 | -99.7949
4.596 | -99.9246
4.5998 | -100.0685
4.6035 | -100.217
4.6073 | -100.3471
4.611 | -100.4469
4.6147 | -100.5057
4.6185 | -100.5171
4.6222 | -100.4907
4.626 | -100.4233
4.6297 | -100.3333
4.6334 | -100.2444
4.6372 | -100.1642
4.6409 | -100.1107
4.6447 | -100.0843
4.6484 | -100.0788
4.6522 | -100.1004
4.6559 | -100.1241
4.6596 | -100.1464
4.6634 | -100.1688
4.6671 | -100.1953
4.6709 | -100.2184
4.6746 | -100.2354
4.6784 | -100.2634
4.6821 | -100.3034
4.6858 | -100.3397
4.6896 | -100.3674
4.6933 | -100.3864
4.6971 | -100.3979
4.7008 | -100.395
4.7046 | -100.3767
4.7083 | -100.355
4.712 | -100.3265
4.7158 | -100.2968
4.7195 | -100.2704
4.7233 | -100.2542
4.727 | -100.2517
4.7308 | -100.2517
4.7345 | -100.2516
4.7382 | -100.2577
4.742 | -100.2662
4.7457 | -100.2766
4.7495 | -100.2874
4.7532 | -100.292
4.757 | -100.3133
4.7607 | -100.3291
4.7644 | -100.33
4.7682 | -100.334
4.7719 | -100.3396
4.7757 | -100.3351
4.7794 | -100.3272
4.7832 | -100.3224
4.7869 | -100.3101
4.7906 | -100.2877
4.7944 | -100.2456
4.7981 | -100.1957
4.8019 | -100.1393
4.8056 | -100.0891
4.8094 | -100.0456
4.8131 | -100.0157
4.8168 | -100.0119
4.8206 | -100.0322
4.8243 | -100.0824
4.8281 | -100.1462
4.8318 | -100.229
4.8356 | -100.3213
4.8393 | -100.3981
4.843 | -100.4571
4.8468 | -100.508
4.8505 | -100.5335
4.8543 | -100.5437
4.858 | -100.5537
4.8618 | -100.5707
4.8655 | -100.5912
4.8692 | -100.6141
4.873 | -100.6308
4.8767 | -100.6428
4.8805 | -100.6445
4.8842 | -100.6385
4.888 | -100.6323
4.8917 | -100.6245
4.8954 | -100.6128
4.8992 | -100.5948
4.9029 | -100.5809
4.9067 | -100.5744
4.9104 | -100.5707
4.9142 | -100.5822
4.9179 | -100.6072
4.9216 | -100.6553
4.9254 | -100.7256
4.9291 | -100.8051
4.9329 | -100.8912
4.9366 | -100.9655
4.9404 | -101.0237
4.9441 | -101.0572
4.9478 | -101.0623
4.9516 | -101.043
4.9553 | -100.9972
4.9591 | -100.94
4.9628 | -100.8842
4.9665 | -100.8305
4.9703 | -100.7851
4.974 | -100.7526
4.9778 | -100.7308
4.9815 | -100.7114
4.9853 | -100.6799
4.989 | -100.6457
4.9927 | -100.6062
4.9965 | -100.5562
5.0002 | -100.498
5.004 | -100.4464
5.0077 | -100.4016
5.0115 | -100.3657
5.0152 | -100.3484
5.0189 | -100.3462
5.0227 | -100.3646
5.0264 | -100.4
5.0302 | -100.4484
5.0339 | -100.5177
5.0377 | -100.5944
5.0414 | -100.6772
5.0451 | -100.7638
5.0489 | -100.8487
5.0526 | -100.9251
5.0564 | -100.9869
5.0601 | -101.0209
5.0639 | -101.0391
5.0676 | -101.0358
5.0713 | -101.0117
5.0751 | -100.9866
5.0788 | -100.955
5.0826 | -100.9185
5.0863 | -100.8888
5.0901 | -100.8605
5.0938 | -100.8359
5.0975 | -100.8212
5.1013 | -100.8202
5.105 | -100.8327
5.1088 | -100.8471
5.1125 | -100.8657
5.1163 | -100.8899
5.12 | -100.9158
5.1237 | -100.9393
5.1275 | -100.9581
5.1312 | -100.9795
5.135 | -100.9983
5.1387 | -101.0172
5.1425 | -101.0169
5.1462 | -100.9929
5.1499 | -100.957
5.1537 | -100.9114
5.1574 | -100.8618
5.1612 | -100.8168
5.1649 | -100.7806
5.1687 | -100.7693
5.1724 | -100.7782
5.1761 | -100.8013
5.1799 | -100.8308
5.1836 | -100.8722
5.1874 | -100.9215
5.1911 | -100.9566
5.1949 | -100.9861
5.1986 | -101.0043
5.2023 | -101.0167
5.2061 | -101.0259
5.2098 | -101.0253
5.2136 | -101.0179
5.2173 | -101.0104
5.2211 | -101.0116
5.2248 | -101.016
5.2285 | -101.0206
5.2323 | -101.0325
5.236 | -101.0554
5.2398 | -101.0808
5.2435 | -101.1049
5.2473 | -101.1105
5.251 | -101.0973
5.2547 | -101.0696
5.2585 | -101.0284
5.2622 | -100.9727
5.266 | -100.9057
5.2697 | -100.8407
5.2735 | -100.7805
5.2772 | -100.7388
5.2809 | -100.7083
5.2847 | -100.6867
5.2884 | -100.6739
5.2922 | -100.6639
5.2959 | -100.6473
5.2996 | -100.62
5.3034 | -100.5843
5.3071 | -100.5381
5.3109 | -100.48
5.3146 | -100.4122
5.3184 | -100.341
5.3221 | -100.2755
5.3258 | -100.2254
5.3296 | -100.2004
5.3333 | -100.194
5.3371 | -100.2134
5.3408 | -100.2569
5.3446 | -100.3276
5.3483 | -100.415
5.352 | -100.5084
5.3558 | -100.6009
5.3595 | -100.6839
5.3633 | -100.7573
5.367 | -100.8167
5.3708 | -100.8564
5.3745 | -100.8732
5.3782 | -100.8736
5.382 | -100.8587
5.3857 | -100.8337
5.3895 | -100.7977
5.3932 | -100.7608
5.397 | -100.7189
5.4007 | -100.6867
5.4044 | -100.6554
5.4082 | -100.632
5.4119 | -100.6177
5.4157 | -100.617
5.4194 | -100.6254
5.4232 | -100.6465
5.4269 | -100.6774
5.4306 | -100.718
5.4344 | -100.7728
5.4381 | -100.8389
5.4419 | -100.914
5.4456 | -100.9892
5.4494 | -101.0669
5.4531 | -101.1453
5.4568 | -101.2194
5.4606 | -101.2873
5.4643 | -101.3548
5.4681 | -101.421
5.4718 | -101.4933
5.4756 | -101.5601
5.4793 | -101.6194
5.483 | -101.6664
5.4868 | -101.6968
5.4905 | -101.7058
5.4943 | -101.69
5.498 | -101.647
5.5018 | -101.5819
5.5055 | -101.5054
5.5092 | -101.4279
5.513 | -101.3537
5.5167 | -101.2939
5.5205 | -101.2471
5.5242 | -101.2196
5.528 | -101.2166
5.5317 | -101.2207
5.5354 | -101.2355
5.5392 | -101.2566
5.5429 | -101.2725
5.5467 | -101.2775
5.5504 | -101.2656
5.5542 | -101.242
5.5579 | -101.2098
5.5616 | -101.1751
5.5654 | -101.1415
5.5691 | -101.1104
5.5729 | -101.0847
5.5766 | -101.0743
5.5804 | -101.0775
5.5841 | -101.0935
5.5878 | -101.1176
5.5916 | -101.147
5.5953 | -101.174
5.5991 | -101.2037
5.6028 | -101.224
5.6066 | -101.231
5.6103 | -101.2211
5.614 | -101.1999
5.6178 | -101.1654
5.6215 | -101.1159
5.6253 | -101.0701
5.629 | -101.0285
5.6327 | -100.9859
5.6365 | -100.9528
5.6402 | -100.9299
5.644 | -100.9119
5.6477 | -100.9087
5.6515 | -100.9196
5.6552 | -100.9448
5.6589 | -100.9801
5.6627 | -101.0206
5.6664 | -101.0681
5.6702 | -101.1172
5.6739 | -101.1625
5.6777 | -101.1847
5.6814 | -101.1834
5.6851 | -101.1593
5.6889 | -101.1204
5.6926 | -101.0683
5.6964 | -101.0019
5.7001 | -100.9255
5.7039 | -100.8502
5.7076 | -100.7772
5.7113 | -100.7115
5.7151 | -100.6691
5.7188 | -100.6554
5.7226 | -100.6694
5.7263 | -100.7075
5.7301 | -100.7756
5.7338 | -100.8642
5.7375 | -100.9632
5.7413 | -101.0667
5.745 | -101.1613
5.7488 | -101.2429
5.7525 | -101.2951
5.7563 | -101.3187
5.76 | -101.3135
5.7637 | -101.2925
5.7675 | -101.2561
5.7712 | -101.204
5.775 | -101.15
5.7787 | -101.1087
5.7825 | -101.0893
5.7862 | -101.09
5.7899 | -101.1106
5.7937 | -101.1411
5.7974 | -101.1796
5.8012 | -101.2173
5.8049 | -101.2484
5.8087 | -101.2664
5.8124 | -101.2724
5.8161 | -101.2652
5.8199 | -101.2506
5.8236 | -101.2328
5.8274 | -101.2196
5.8311 | -101.2214
5.8349 | -101.2336
5.8386 | -101.2632
5.8423 | -101.3103
5.8461 | -101.3769
5.8498 | -101.4503
5.8536 | -101.5278
5.8573 | -101.6143
5.8611 | -101.6956
5.8648 | -101.7564
5.8685 | -101.7978
5.8723 | -101.8269
5.876 | -101.8384
5.8798 | -101.83
5.8835 | -101.8032
5.8873 | -101.7655
5.891 | -101.7137
5.8947 | -101.6609
5.8985 | -101.6108
5.9022 | -101.5575
5.906 | -101.5042
5.9097 | -101.4607
5.9135 | -101.4287
5.9172 | -101.4116
5.9209 | -101.4181
5.9247 | -101.4462
5.9284 | -101.4807
5.9322 | -101.5267
5.9359 | -101.5841
5.9396 | -101.6301
5.9434 | -101.6728
5.9471 | -101.7036
5.9509 | -101.7223
5.9546 | -101.7221
5.9584 | -101.7134
5.9621 | -101.6948
5.9658 | -101.6677
5.9696 | -101.638
5.9733 | -101.6046
5.9771 | -101.5688
5.9808 | -101.5324
5.9846 | -101.4877
5.9883 | -101.4461
5.992 | -101.4086
5.9958 | -101.3715
5.9995 | -101.3354
6.0033 | -101.2993
6.007 | -101.2682
6.0108 | -101.2407
6.0145 | -101.2211
6.0182 | -101.2132
6.022 | -101.2102
6.0257 | -101.2195
6.0295 | -101.2423
6.0332 | -101.2841
6.037 | -101.338
6.0407 | -101.3983
6.0444 | -101.4662
6.0482 | -101.5385
6.0519 | -101.5953
6.0557 | -101.6308
6.0594 | -101.6489
6.0632 | -101.6426
6.0669 | -101.6117
6.0706 | -101.562
6.0744 | -101.4945
6.0781 | -101.4183
6.0819 | -101.3368
6.0856 | -101.2608
6.0894 | -101.202
6.0931 | -101.1627
6.0968 | -101.1425
6.1006 | -101.1418
6.1043 | -101.1658
6.1081 | -101.2083
6.1118 | -101.2703
6.1156 | -101.3423
6.1193 | -101.4211
6.123 | -101.4984
6.1268 | -101.5736
6.1305 | -101.6461
6.1343 | -101.716
6.138 | -101.778
6.1418 | -101.831
6.1455 | -101.8626
6.1492 | -101.8795
6.153 | -101.8842
6.1567 | -101.8695
6.1605 | -101.8355
6.1642 | -101.7844
6.168 | -101.7253
6.1717 | -101.661
6.1754 | -101.594
6.1792 | -101.5281
6.1829 | -101.4669
6.1867 | -101.4169
6.1904 | -101.3824
6.1942 | -101.358
6.1979 | -101.3459
6.2016 | -101.335
6.2054 | -101.3267
6.2091 | -101.3198
6.2129 | -101.3091
6.2166 | -101.3033
6.2204 | -101.2966
6.2241 | -101.2925
6.2278 | -101.2883
6.2316 | -101.2915
6.2353 | -101.3052
6.2391 | -101.3342
6.2428 | -101.3673
6.2466 | -101.402
6.2503 | -101.4453
6.254 | -101.5009
6.2578 | -101.5632
6.2615 | -101.6264
6.2653 | -101.6814
6.269 | -101.7318
6.2727 | -101.7763
6.2765 | -101.7992
6.2802 | -101.804
6.284 | -101.7892
6.2877 | -101.7617
6.2915 | -101.7196
6.2952 | -101.6745
6.2989 | -101.6274
6.3027 | -101.5836
6.3064 | -101.5494
6.3102 | -101.5327
6.3139 | -101.5381
6.3177 | -101.5637
6.3214 | -101.6147
6.3251 | -101.6866
6.3289 | -101.7746
6.3326 | -101.8746
6.3364 | -101.9813
6.3401 | -102.0894
6.3439 | -102.1932
6.3476 | -102.2753
6.3513 | -102.3373
6.3551 | -102.3824
6.3588 | -102.4137
6.3626 | -102.423
6.3663 | -102.4147
6.3701 | -102.3993
6.3738 | -102.3736
6.3775 | -102.3425
6.3813 | -102.309
6.385 | -102.2785
6.3888 | -102.2489
6.3925 | -102.2224
6.3963 | -102.2007
6.4 | -102.1863
6.4037 | -102.1772
6.4075 | -102.1755
6.4112 | -102.1821
6.415 | -102.1942
6.4187 | -102.2043
6.4225 | -102.2151
6.4262 | -102.2305
6.4299 | -102.2494
6.4337 | -102.2671
6.4374 | -102.2923
6.4412 | -102.33
6.4449 | -102.3685
6.4487 | -102.4068
6.4524 | -102.4471
6.4561 | -102.4899
6.4599 | -102.5293
6.4636 | -102.5597
6.4674 | -102.5835
6.4711 | -102.6039
6.4749 | -102.6193
6.4786 | -102.6396
6.4823 | -102.66
6.4861 | -102.6836
6.4898 | -102.7096
6.4936 | -102.7373
6.4973 | -102.7607
6.5011 | -102.7814
6.5048 | -102.7968
6.5085 | -102.8037
6.5123 | -102.805
6.516 | -102.7985
6.5198 | -102.7795
6.5235 | -102.7472
6.5273 | -102.7105
6.531 | -102.6667
6.5347 | -102.6206
6.5385 | -102.5725
6.5422 | -102.5324
6.546 | -102.4916
6.5497 | -102.4521
6.5534 | -102.4221
6.5572 | -102.3982
6.5609 | -102.3823
6.5647 | -102.3741
6.5684 | -102.3639
6.5722 | -102.3539
6.5759 | -102.3476
6.5796 | -102.3399
6.5834 | -102.3302
6.5871 | -102.3228
6.5909 | -102.3101
6.5946 | -102.2982
6.5984 | -102.2913
6.6021 | -102.279
6.6058 | -102.2625
6.6096 | -102.2439
6.6133 | -102.2223
6.6171 | -102.2057
6.6208 | -102.1917
6.6246 | -102.1798
6.6283 | -102.1677
6.632 | -102.1617
6.6358 | -102.1574
6.6395 | -102.1612
6.6433 | -102.1642
6.647 | -102.1772
6.6508 | -102.1918
6.6545 | -102.2028
6.6582 | -102.221
6.662 | -102.2405
6.6657 | -102.2544
6.6695 | -102.2702
6.6732 | -102.2837
6.677 | -102.2966
6.6807 | -102.3112
6.6844 | -102.3218
6.6882 | -102.3298
6.6919 | -102.3408
6.6957 | -102.3508
6.6994 | -102.3551
6.7032 | -102.3613
6.7069 | -102.3763
6.7106 | -102.3972
6.7144 | -102.4181
6.7181 | -102.4344
6.7219 | -102.4546
6.7256 | -102.4772
6.7294 | -102.502
6.7331 | -102.5266
6.7368 | -102.5498
6.7406 | -102.5665
6.7443 | -102.5846
6.7481 | -102.5902
6.7518 | -102.5904
6.7556 | -102.5861
6.7593 | -102.5726
6.763 | -102.5497
6.7668 | -102.5252
6.7705 | -102.5005
6.7743 | -102.469
6.778 | -102.4321
6.7818 | -102.3974
6.7855 | -102.3746
6.7892 | -102.3542
6.793 | -102.3363
6.7967 | -102.3261
6.8005 | -102.3281
6.8042 | -102.3387
6.808 | -102.3482
6.8117 | -102.3612
6.8154 | -102.388
6.8192 | -102.4228
6.8229 | -102.4549
6.8267 | -102.4851
6.8304 | -102.515
6.8342 | -102.5438
6.8379 | -102.5682
6.8416 | -102.5952
6.8454 | -102.6185
6.8491 | -102.6411
6.8529 | -102.6651
6.8566 | -102.688
6.8604 | -102.7074
6.8641 | -102.7288
6.8678 | -102.7482
6.8716 | -102.762
6.8753 | -102.7719
6.8791 | -102.7777
6.8828 | -102.7842
6.8865 | -102.7884
6.8903 | -102.7981
6.894 | -102.7986
6.8978 | -102.7944
6.9015 | -102.7902
6.9053 | -102.784
6.909 | -102.7639
6.9127 | -102.7291
6.9165 | -102.688
6.9202 | -102.6445
6.924 | -102.5891
6.9277 | -102.5269
6.9315 | -102.4612
6.9352 | -102.3957
6.9389 | -102.3358
6.9427 | -102.2768
6.9464 | -102.2266
6.9502 | -102.1885
6.9539 | -102.165
6.9577 | -102.1638
6.9614 | -102.1845
6.9651 | -102.2212
6.9689 | -102.2732
6.9726 | -102.3398
6.9764 | -102.4197
6.9801 | -102.501
6.9839 | -102.5804
6.9876 | -102.6587
6.9913 | -102.7317
6.9951 | -102.8085
6.9988 | -102.8757
7.0026 | -102.9252
7.0063 | -102.9645
7.0101 | -103.004
7.0138 | -103.0288
7.0175 | -103.0034
7.0213 | -102.9951
7.025 | -102.9687
7.0288 | -102.9444
7.0325 | -102.9158
7.0363 | -102.8804
7.04 | -102.8407
7.0437 | -102.8045
7.0475 | -102.7763
7.0512 | -102.7561
7.055 | -102.7398
7.0587 | -102.7312
7.0625 | -102.7315
7.0662 | -102.74
7.0699 | -102.7577
7.0737 | -102.78
7.0774 | -102.8087
7.0812 | -102.8395
7.0849 | -102.8756
7.0887 | -102.9088
7.0924 | -102.9384
7.0961 | -102.9632
7.0999 | -102.9763
7.1036 | -102.9815
7.1074 | -102.9781
7.1111 | -102.9613
7.1149 | -102.9295
7.1186 | -102.8894
7.1223 | -102.8432
7.1261 | -102.7922
7.1298 | -102.7362
7.1336 | -102.685
7.1373 | -102.6251
7.1411 | -102.5683
7.1448 | -102.5235
7.1485 | -102.4912
7.1523 | -102.4665
7.156 | -102.4548
7.1598 | -102.4574
7.1635 | -102.4791
7.1673 | -102.5104
7.171 | -102.55
7.1747 | -102.5997
7.1785 | -102.6582
7.1822 | -102.7229
7.186 | -102.7871
7.1897 | -102.8546
7.1935 | -102.9157
7.1972 | -102.9752
7.2009 | -103.0249
7.2047 | -103.0734
7.2084 | -103.1171
7.2122 | -103.1519
7.2159 | -103.1767
7.2196 | -103.1906
7.2234 | -103.1961
7.2271 | -103.193
7.2309 | -103.1733
7.2346 | -103.1473
7.2384 | -103.1107
7.2421 | -103.0661
7.2458 | -103.0258
7.2496 | -102.9801
7.2533 | -102.9339
7.2571 | -102.8917
7.2608 | -102.8626
7.2646 | -102.8455
7.2683 | -102.8396
7.272 | -102.842
7.2758 | -102.8542
7.2795 | -102.8705
7.2833 | -102.8894
7.287 | -102.9058
7.2908 | -102.9209
7.2945 | -102.9334
7.2982 | -102.9383
7.302 | -102.9371
7.3057 | -102.9265
7.3095 | -102.9134
7.3132 | -102.8977
7.317 | -102.8802
7.3207 | -102.866
7.3244 | -102.8543
7.3282 | -102.8395
7.3319 | -102.833
7.3357 | -102.8271
7.3394 | -102.8267
7.3432 | -102.8224
7.3469 | -102.8239
7.3506 | -102.8278
7.3544 | -102.8302
7.3581 | -102.8337
7.3619 | -102.8364
7.3656 | -102.8346
7.3694 | -102.8244
7.3731 | -102.8156
7.3768 | -102.8143
7.3806 | -102.8153
7.3843 | -102.8208
7.3881 | -102.8356
7.3918 | -102.8528
7.3956 | -102.8672
7.3993 | -102.8687
7.403 | -102.8682
7.4068 | -102.8688
7.4105 | -102.8701
7.4143 | -102.8719
7.418 | -102.8716
7.4218 | -102.8683
7.4255 | -102.8729
7.4292 | -102.8797
7.433 | -102.8906
7.4367 | -102.915
7.4405 | -102.9346
7.4442 | -102.9576
7.448 | -102.9826
7.4517 | -103.0169
7.4554 | -103.0502
7.4592 | -103.08
7.4629 | -103.1078
7.4667 | -103.1338
7.4704 | -103.1581
7.4742 | -103.1834
7.4779 | -103.2079
7.4816 | -103.2163
7.4854 | -103.2096
7.4891 | -103.2008
7.4929 | -103.1864
7.4966 | -103.1633
7.5004 | -103.1372
7.5041 | -103.1112
7.5078 | -103.0802
7.5116 | -103.038
7.5153 | -102.9839
7.5191 | -102.925
7.5228 | -102.8624
7.5266 | -102.7956
7.5303 | -102.7316
7.534 | -102.6715
7.5378 | -102.6166
7.5415 | -102.5699
7.5453 | -102.5253
7.549 | -102.4931
7.5527 | -102.4759
7.5565 | -102.504
7.5602 | -102.5114
7.564 | -102.5295
7.5677 | -102.5584
7.5715 | -102.6034
7.5752 | -102.6576
7.5789 | -102.717
7.5827 | -102.7815
7.5864 | -102.852
7.5902 | -102.9173
7.5939 | -102.979
7.5977 | -103.0388
7.6014 | -103.0867
7.6051 | -103.1201
7.6089 | -103.1418
7.6126 | -103.157
7.6164 | -103.164
7.6201 | -103.1626
7.6239 | -103.1499
7.6276 | -103.1318
7.6313 | -103.1117
7.6351 | -103.0909
7.6388 | -103.0728
7.6426 | -103.0533
7.6463 | -103.0347
7.6501 | -103.0161
7.6538 | -103.003
7.6575 | -102.9985
7.6613 | -102.9937
7.665 | -102.9892
7.6688 | -102.9875
7.6725 | -102.9942
7.6763 | -103.0068
7.68 | -103.021
7.6837 | -103.0374
7.6875 | -103.0577
7.6912 | -103.0796
7.695 | -103.1015
7.6987 | -103.1253
7.7025 | -103.1446
7.7062 | -103.1736
7.7099 | -103.2043
7.7137 | -103.228
7.7174 | -103.2478
7.7212 | -103.2638
7.7249 | -103.2737
7.7287 | -103.2801
7.7324 | -103.2834
7.7361 | -103.2882
7.7399 | -103.2909
7.7436 | -103.2889
7.7474 | -103.2762
7.7511 | -103.2642
7.7549 | -103.2541
7.7586 | -103.2435
7.7623 | -103.2347
7.7661 | -103.2276
7.7698 | -103.22
7.7736 | -103.2106
7.7773 | -103.203
7.7811 | -103.1983
7.7848 | -103.1962
7.7885 | -103.1942
7.7923 | -103.1931
7.796 | -103.1924
7.7998 | -103.1976
7.8035 | -103.2009
7.8073 | -103.2086
7.811 | -103.2213
7.8147 | -103.232
7.8185 | -103.246
7.8222 | -103.2583
7.826 | -103.2747
7.8297 | -103.2969
7.8335 | -103.3246
7.8372 | -103.3579
7.8409 | -103.393
7.8447 | -103.4285
7.8484 | -103.4707
7.8522 | -103.5147
7.8559 | -103.5592
7.8596 | -103.6084
7.8634 | -103.6626
7.8671 | -103.7134
7.8709 | -103.7638
7.8746 | -103.8115
7.8784 | -103.8623
7.8821 | -103.9164
7.8858 | -103.9651
7.8896 | -104.0048
7.8933 | -104.0407
7.8971 | -104.0641
7.9008 | -104.0785
7.9046 | -104.0787
7.9083 | -104.0683
7.912 | -104.053
7.9158 | -104.0327
7.9195 | -103.9973
7.9233 | -103.949
7.927 | -103.8983
7.9308 | -103.8504
7.9345 | -103.7974
7.9382 | -103.7331
7.942 | -103.6744
7.9457 | -103.6186
7.9495 | -103.5604
7.9532 | -103.5046
7.957 | -103.4518
7.9607 | -103.4044
7.9644 | -103.3604
7.9682 | -103.3237
7.9719 | -103.3
7.9757 | -103.2806
7.9794 | -103.2669
7.9832 | -103.2614
7.9869 | -103.2595
7.9906 | -103.2598
7.9944 | -103.2616
7.9981 | -103.2627
8.0019 | -103.2677
8.0056 | -103.2751
8.0094 | -103.2764
8.0131 | -103.2724
8.0168 | -103.2643
8.0206 | -103.2554
8.0243 | -103.2412
8.0281 | -103.2203
8.0318 | -103.1981
8.0356 | -103.1862
8.0393 | -103.1803
8.043 | -103.1728
8.0468 | -103.1687
8.0505 | -103.1755
8.0543 | -103.1838
8.058 | -103.185
8.0618 | -103.185
8.0655 | -103.1891
8.0692 | -103.1916
8.073 | -103.1918
8.0767 | -103.1842
8.0805 | -103.1758
8.0842 | -103.1722
8.088 | -103.1675
8.0917 | -103.1701
8.0954 | -103.1773
8.0992 | -103.204
8.1029 | -103.2214
8.1067 | -103.2457
8.1104 | -103.2801
8.1142 | -103.3127
8.1179 | -103.3417
8.1216 | -103.3659
8.1254 | -103.3894
8.1291 | -103.4081
8.1329 | -103.4295
8.1366 | -103.4451
8.1404 | -103.4537
8.1441 | -103.4571
8.1478 | -103.4585
8.1516 | -103.4516
8.1553 | -103.4357
8.1591 | -103.4166
8.1628 | -103.4001
8.1666 | -103.3829
8.1703 | -103.3673
8.174 | -103.3631
8.1778 | -103.3654
8.1815 | -103.3716
8.1853 | -103.3839
8.189 | -103.3994
8.1927 | -103.4233
8.1965 | -103.4549
8.2002 | -103.4912
8.204 | -103.5282
8.2077 | -103.5652
8.2115 | -103.6028
8.2152 | -103.6405
8.2189 | -103.6705
8.2227 | -103.6889
8.2264 | -103.6988
8.2302 | -103.7034
8.2339 | -103.6923
8.2377 | -103.6699
8.2414 | -103.6459
8.2451 | -103.6158
8.2489 | -103.5744
8.2526 | -103.5299
8.2564 | -103.4792
8.2601 | -103.4256
8.2639 | -103.3763
8.2676 | -103.3339
8.2713 | -103.2908
8.2751 | -103.2512
8.2788 | -103.2193
8.2826 | -103.1965
8.2863 | -103.1749
8.2901 | -103.1594
8.2938 | -103.1556
8.2975 | -103.1622
8.3013 | -103.1742
8.305 | -103.1951
8.3088 | -103.2235
8.3125 | -103.2597
8.3163 | -103.3033
8.32 | -103.3486
8.3237 | -103.3998
8.3275 | -103.4506
8.3312 | -103.5018
8.335 | -103.5522
8.3387 | -103.6032
8.3425 | -103.6517
8.3462 | -103.6955
8.3499 | -103.7343
8.3537 | -103.7651
8.3574 | -103.7936
8.3612 | -103.82
8.3649 | -103.8371
8.3687 | -103.8497
8.3724 | -103.8615
8.3761 | -103.8689
8.3799 | -103.87
8.3836 | -103.8731
8.3874 | -103.8733
8.3911 | -103.8735
8.3949 | -103.8804
8.3986 | -103.8969
8.4023 | -103.9139
8.4061 | -103.9276
8.4098 | -103.9493
8.4136 | -103.9752
8.4173 | -103.9973
8.4211 | -104.0127
8.4248 | -104.023
8.4285 | -104.0225
8.4323 | -104.0132
8.436 | -103.9973
8.4398 | -103.9677
8.4435 | -103.9333
8.4473 | -103.8827
8.451 | -103.8264
8.4547 | -103.7716
8.4585 | -103.7138
8.4622 | -103.6561
8.466 | -103.6043
8.4697 | -103.5606
8.4734 | -103.5235
8.4772 | -103.4909
8.4809 | -103.4711
8.4847 | -103.4539
8.4884 | -103.447
8.4922 | -103.4456
8.4959 | -103.4523
8.4996 | -103.4665
8.5034 | -103.4829
8.5071 | -103.5012
8.5109 | -103.523
8.5146 | -103.5464
8.5184 | -103.5677
8.5221 | -103.585
8.5258 | -103.6029
8.5296 | -103.6202
8.5333 | -103.6399
8.5371 | -103.6556
8.5408 | -103.6729
8.5446 | -103.6908
8.5483 | -103.7109
8.552 | -103.7362
8.5558 | -103.7573
8.5595 | -103.7759
8.5633 | -103.7955
8.567 | -103.8123
8.5708 | -103.8327
8.5745 | -103.8527
8.5782 | -103.868
8.582 | -103.8823
8.5857 | -103.895
8.5895 | -103.9072
8.5932 | -103.9184
8.597 | -103.9255
8.6007 | -103.9382
8.6044 | -103.9463
8.6082 | -103.9545
8.6119 | -103.958
8.6157 | -103.9594
8.6194 | -103.9608
8.6232 | -103.9618
8.6269 | -103.9638
8.6306 | -103.952
8.6344 | -103.9303
8.6381 | -103.8861
8.6419 | -103.8577
8.6456 | -103.8231
8.6494 | -103.7842
8.6531 | -103.7423
8.6568 | -103.6976
8.6606 | -103.6398
8.6643 | -103.5819
8.6681 | -103.521
8.6718 | -103.4672
8.6756 | -103.4203
8.6793 | -103.3766
8.683 | -103.3408
8.6868 | -103.3128
8.6905 | -103.2937
8.6943 | -103.2857
8.698 | -103.2868
8.7018 | -103.2987
8.7055 | -103.3211
8.7092 | -103.3483
8.713 | -103.3813
8.7167 | -103.42
8.7205 | -103.4623
8.7242 | -103.5057
8.728 | -103.5431
8.7317 | -103.5859
8.7354 | -103.6298
8.7392 | -103.6703
8.7429 | -103.7059
8.7467 | -103.7353
8.7504 | -103.7661
8.7542 | -103.7963
8.7579 | -103.8307
8.7616 | -103.8621
8.7654 | -103.8829
8.7691 | -103.9005
8.7729 | -103.9197
8.7766 | -103.9428
8.7804 | -103.9653
8.7841 | -103.9868
8.7878 | -104.0065
8.7916 | -104.0261
8.7953 | -104.0388
8.7991 | -104.0439
8.8028 | -104.0413
8.8065 | -104.0329
8.8103 | -104.0217
8.814 | -104.0056
8.8178 | -103.9862
8.8215 | -103.9612
8.8253 | -103.9308
8.829 | -103.8964
8.8327 | -103.8566
8.8365 | -103.8142
8.8402 | -103.7772
8.844 | -103.7401
8.8477 | -103.6948
8.8515 | -103.6535
8.8552 | -103.6227
8.8589 | -103.5977
8.8627 | -103.5717
8.8664 | -103.5515
8.8702 | -103.5388
8.8739 | -103.5343
8.8777 | -103.5343
8.8814 | -103.535
8.8851 | -103.5471
8.8889 | -103.5669
8.8926 | -103.5934
8.8964 | -103.6273
8.9001 | -103.6621
8.9039 | -103.697
8.9076 | -103.7377
8.9113 | -103.7813
8.9151 | -103.8292
8.9188 | -103.8808
8.9226 | -103.9347
8.9263 | -103.9903
8.9301 | -104.0471
8.9338 | -104.0993
8.9375 | -104.1487
8.9413 | -104.1939
8.945 | -104.2388
8.9488 | -104.2736
8.9525 | -104.2998
8.9563 | -104.317
8.96 | -104.3311
8.9637 | -104.3355
8.9675 | -104.3334
8.9712 | -104.323
8.975 | -104.3019
8.9787 | -104.2822
8.9825 | -104.2507
8.9862 | -104.2125
8.9899 | -104.1742
8.9937 | -104.1375
8.9974 | -104.1037
9.0012 | -104.0749
9.0049 | -104.0504
9.0087 | -104.032
9.0124 | -104.021
9.0161 | -104.0101
9.0199 | -104.0049
9.0236 | -104.0065
9.0274 | -104.0165
9.0311 | -104.0306
9.0349 | -104.0519
9.0386 | -104.0753
9.0423 | -104.1068
9.0461 | -104.1432
9.0498 | -104.1818
9.0536 | -104.223
9.0573 | -104.2666
9.0611 | -104.3109
9.0648 | -104.3569
9.0685 | -104.3996
9.0723 | -104.4391
9.076 | -104.4679
9.0798 | -104.4936
9.0835 | -104.5131
9.0873 | -104.5227
9.091 | -104.5148
9.0947 | -104.5018
9.0985 | -104.4752
9.1022 | -104.4397
9.106 | -104.4001
9.1097 | -104.3561
9.1135 | -104.3106
9.1172 | -104.2603
9.1209 | -104.2035
9.1247 | -104.1438
9.1284 | -104.0889
9.1322 | -104.0382
9.1359 | -103.9926
9.1396 | -103.9509
9.1434 | -103.9188
9.1471 | -103.8951
9.1509 | -103.8801
9.1546 | -103.8751
9.1584 | -103.8756
9.1621 | -103.8846
9.1658 | -103.8985
9.1696 | -103.9174
9.1733 | -103.9378
9.1771 | -103.9657
9.1808 | -103.9876
9.1846 | -104.0144
9.1883 | -104.041
9.192 | -104.0655
9.1958 | -104.087
9.1995 | -104.1059
9.2033 | -104.1249
9.207 | -104.1449
9.2108 | -104.1632
9.2145 | -104.1782
9.2182 | -104.1893
9.222 | -104.1998
9.2257 | -104.2134
9.2295 | -104.2267
9.2332 | -104.2431
9.237 | -104.2603
9.2407 | -104.2765
9.2444 | -104.2944
9.2482 | -104.3137
9.2519 | -104.3298
9.2557 | -104.3529
9.2594 | -104.3809
9.2632 | -104.4025
9.2669 | -104.4202
9.2706 | -104.4351
9.2744 | -104.4534
9.2781 | -104.466
9.2819 | -104.4716
9.2856 | -104.4675
9.2894 | -104.46
9.2931 | -104.4491
9.2968 | -104.4332
9.3006 | -104.418
9.3043 | -104.4025
9.3081 | -104.3826
9.3118 | -104.361
9.3156 | -104.3432
9.3193 | -104.3295
9.323 | -104.3161
9.3268 | -104.3034
9.3305 | -104.293
9.3343 | -104.285
9.338 | -104.2757
9.3418 | -104.2678
9.3455 | -104.2641
9.3492 | -104.2659
9.353 | -104.264
9.3567 | -104.2607
9.3605 | -104.2596
9.3642 | -104.2569
9.368 | -104.2493
9.3717 | -104.24
9.3754 | -104.2304
9.3792 | -104.2195
9.3829 | -104.1991
9.3867 | -104.1739
9.3904 | -104.1548
9.3942 | -104.1325
9.3979 | -104.102
9.4016 | -104.0756
9.4054 | -104.0524
9.4091 | -104.0343
9.4129 | -104.0183
9.4166 | -104.0045
9.4204 | -104.0036
9.4241 | -104.0116
9.4278 | -104.0199
9.4316 | -104.0346
9.4353 | -104.0542
9.4391 | -104.0765
9.4428 | -104.1037
9.4466 | -104.1324
9.4503 | -104.1601
9.454 | -104.1838
9.4578 | -104.2061
9.4615 | -104.2318
9.4653 | -104.2586
9.469 | -104.2818
9.4727 | -104.3019
9.4765 | -104.3248
9.4802 | -104.3448
9.484 | -104.3632
9.4877 | -104.3812
9.4915 | -104.4063
9.4952 | -104.4303
9.4989 | -104.4522
9.5027 | -104.4802
9.5064 | -104.5135
9.5102 | -104.5439
9.5139 | -104.5763
9.5177 | -104.6121
9.5214 | -104.6491
9.5251 | -104.684
9.5289 | -104.713
9.5326 | -104.7412
9.5364 | -104.7705
9.5401 | -104.7908
9.5439 | -104.7986
9.5476 | -104.7995
9.5513 | -104.7936
9.5551 | -104.7857
9.5588 | -104.7676
9.5626 | -104.7424
9.5663 | -104.7137
9.5701 | -104.6816
9.5738 | -104.644
9.5775 | -104.5998
9.5813 | -104.5552
9.585 | -104.5106
9.5888 | -104.4681
9.5925 | -104.4239
9.5963 | -104.378
9.6 | -104.3276
9.6037 | -104.2827
9.6075 | -104.2426
9.6112 | -104.2077
9.615 | -104.1749
9.6187 | -104.1513
9.6225 | -104.1352
9.6262 | -104.1239
9.6299 | -104.1178
9.6337 | -104.1212
9.6374 | -104.1339
9.6412 | -104.1531
9.6449 | -104.1796
9.6487 | -104.2094
9.6524 | -104.2387
9.6561 | -104.2768
9.6599 | -104.3216
9.6636 | -104.369
9.6674 | -104.4157
9.6711 | -104.4635
9.6749 | -104.5088
9.6786 | -104.5523
9.6823 | -104.5976
9.6861 | -104.6372
9.6898 | -104.6706
9.6936 | -104.6935
9.6973 | -104.7154
9.7011 | -104.7334
9.7048 | -104.7429
9.7085 | -104.7405
9.7123 | -104.7317
9.716 | -104.7066
9.7198 | -104.6943
9.7235 | -104.6739
9.7273 | -104.6527
9.731 | -104.6244
9.7347 | -104.5993
9.7385 | -104.5786
9.7422 | -104.5598
9.746 | -104.5411
9.7497 | -104.5229
9.7535 | -104.5079
9.7572 | -104.498
9.7609 | -104.4905
9.7647 | -104.4843
9.7684 | -104.4846
9.7722 | -104.4821
9.7759 | -104.4837
9.7796 | -104.4838
9.7834 | -104.4806
9.7871 | -104.4743
9.7909 | -104.473
9.7946 | -104.4683
9.7984 | -104.4486
9.8021 | -104.4252
9.8058 | -104.3994
9.8096 | -104.3684
9.8133 | -104.3303
9.8171 | -104.29
9.8208 | -104.2483
9.8246 | -104.2045
9.8283 | -104.1581
9.832 | -104.1142
9.8358 | -104.0709
9.8395 | -104.024
9.8433 | -103.9828
9.847 | -103.9514
9.8508 | -103.93
9.8545 | -103.9176
9.8582 | -103.9098
9.862 | -103.9088
9.8657 | -103.9164
9.8695 | -103.9342
9.8732 | -103.9565
9.877 | -103.9859
9.8807 | -104.0183
9.8844 | -104.0535
9.8882 | -104.0921
9.8919 | -104.1298
9.8957 | -104.1635
9.8994 | -104.1957
9.9032 | -104.2281
9.9069 | -104.2593
9.9106 | -104.2913
9.9144 | -104.3183
9.9181 | -104.3473
9.9219 | -104.3722
9.9256 | -104.3951
9.9294 | -104.4192
9.9331 | -104.4447
9.9368 | -104.4699
9.9406 | -104.4943
9.9443 | -104.5201
9.9481 | -104.5458
9.9518 | -104.571
9.9556 | -104.5969
9.9593 | -104.6234
9.963 | -104.6473
9.9668 | -104.6711
9.9705 | -104.6927
9.9743 | -104.7136
9.978 | -104.7261
9.9818 | -104.738
9.9855 | -104.7501
9.9892 | -104.7562
9.993 | -104.7594
9.9967 | -104.757
10.0005 | -104.7522
10.0042 | -104.7472
10.008 | -104.7366
10.0117 | -104.7216
10.0154 | -104.7059
10.0192 | -104.6889
10.0229 | -104.6686
10.0267 | -104.6427
10.0304 | -104.6174
10.0342 | -104.5971
10.0379 | -104.5783
10.0416 | -104.5674
10.0454 | -104.5592
10.0491 | -104.5535
10.0529 | -104.5521
10.0566 | -104.5583
10.0604 | -104.5709
10.0641 | -104.586
10.0678 | -104.606
10.0716 | -104.6294
10.0753 | -104.6583
10.0791 | -104.6893
10.0828 | -104.7183
10.0866 | -104.7507
10.0903 | -104.7865
10.094 | -104.8209
10.0978 | -104.8517
10.1015 | -104.8777
10.1053 | -104.9009
10.109 | -104.9245
10.1127 | -104.9424
10.1165 | -104.9565
10.1202 | -104.9651
10.124 | -104.9744
10.1277 | -104.9785
10.1315 | -104.9724
10.1352 | -104.9605
10.1389 | -104.9401
10.1427 | -104.9148
10.1464 | -104.8864
10.1502 | -104.8535
10.1539 | -104.8232
10.1577 | -104.7897
10.1614 | -104.7524
10.1651 | -104.7172
10.1689 | -104.6818
10.1726 | -104.6524
10.1764 | -104.6234
10.1801 | -104.5942
10.1839 | -104.569
10.1876 | -104.5422
10.1913 | -104.5181
10.1951 | -104.4944
10.1988 | -104.4714
10.2026 | -104.4526
10.2063 | -104.4391
10.2101 | -104.4287
10.2138 | -104.4196
10.2175 | -104.4137
10.2213 | -104.405
10.225 | -104.3997
10.2288 | -104.3984
10.2325 | -104.3991
10.2363 | -104.4019
10.24 | -104.4065
10.2437 | -104.4178
10.2475 | -104.4347
10.2512 | -104.4521
10.255 | -104.4687
10.2587 | -104.4906
10.2625 | -104.5047
10.2662 | -104.5177
10.2699 | -104.5321
10.2737 | -104.5455
10.2774 | -104.5562
10.2812 | -104.5671
10.2849 | -104.5756
10.2887 | -104.5802
10.2924 | -104.5852
10.2961 | -104.5864
10.2999 | -104.5874
10.3036 | -104.5823
10.3074 | -104.5733
10.3111 | -104.5589
10.3149 | -104.5452
10.3186 | -104.5295
10.3223 | -104.5105
10.3261 | -104.4907
10.3298 | -104.4725
10.3336 | -104.4579
10.3373 | -104.4488
10.3411 | -104.4421
10.3448 | -104.4417
10.3485 | -104.4462
10.3523 | -104.4511
10.356 | -104.4635
10.3598 | -104.4819
10.3635 | -104.5087
10.3673 | -104.5439
10.371 | -104.5897
10.3747 | -104.6374
10.3785 | -104.6889
10.3822 | -104.743
10.386 | -104.7975
10.3897 | -104.853
10.3934 | -104.9068
10.3972 | -104.9531
10.4009 | -104.999
10.4047 | -105.042
10.4084 | -105.0827
10.4122 | -105.1124
10.4159 | -105.1336
10.4196 | -105.1489
10.4234 | -105.1548
10.4271 | -105.1535
10.4309 | -105.1465
10.4346 | -105.1305
10.4384 | -105.1102
10.4421 | -105.0876
10.4458 | -105.0557
10.4496 | -105.017
10.4533 | -104.9755
10.4571 | -104.9333
10.4608 | -104.8935
10.4646 | -104.8574
10.4683 | -104.8189
10.472 | -104.781
10.4758 | -104.7452
10.4795 | -104.7133
10.4833 | -104.6823
10.487 | -104.6483
10.4908 | -104.6147
10.4945 | -104.5848
10.4982 | -104.5576
10.502 | -104.5356
10.5057 | -104.5205
10.5095 | -104.5045
10.5132 | -104.4903
10.517 | -104.4771
10.5207 | -104.4662
10.5244 | -104.4555
10.5282 | -104.4491
10.5319 | -104.4434
10.5357 | -104.4408
10.5394 | -104.4421
10.5432 | -104.4436
10.5469 | -104.446
10.5506 | -104.4529
10.5544 | -104.455
10.5581 | -104.4582
10.5619 | -104.4582
10.5656 | -104.4577
10.5694 | -104.4603
10.5731 | -104.4579
10.5768 | -104.4477
10.5806 | -104.4379
10.5843 | -104.4311
10.5881 | -104.4247
10.5918 | -104.4177
10.5956 | -104.4113
10.5993 | -104.4092
10.603 | -104.4059
10.6068 | -104.4001
10.6105 | -104.3971
10.6143 | -104.3899
10.618 | -104.3782
10.6218 | -104.3668
10.6255 | -104.3583
10.6292 | -104.3492
10.633 | -104.3387
10.6367 | -104.3316
10.6405 | -104.3271
10.6442 | -104.3211
10.648 | -104.3136
10.6517 | -104.3056
10.6554 | -104.2964
10.6592 | -104.2897
10.6629 | -104.2852
10.6667 | -104.28
10.6704 | -104.2747
10.6742 | -104.2724
10.6779 | -104.2761
10.6816 | -104.283
10.6854 | -104.2899
10.6891 | -104.2969
10.6929 | -104.3056
10.6966 | -104.3145
10.7004 | -104.3254
10.7041 | -104.338
10.7078 | -104.356
10.7116 | -104.3705
10.7153 | -104.3846
10.7191 | -104.4028
10.7228 | -104.4172
10.7265 | -104.4278
10.7303 | -104.44
10.734 | -104.4459
10.7378 | -104.4435
10.7415 | -104.4362
10.7453 | -104.4292
10.749 | -104.4138
10.7527 | -104.399
10.7565 | -104.3881
10.7602 | -104.3772
10.764 | -104.3614
10.7677 | -104.3431
10.7715 | -104.3234
10.7752 | -104.3013
10.7789 | -104.2778
10.7827 | -104.2577
10.7864 | -104.2371
10.7902 | -104.2096
10.7939 | -104.1837
10.7977 | -104.169
10.8014 | -104.1516
10.8051 | -104.1369
10.8089 | -104.1247
10.8126 | -104.1159
10.8164 | -104.108
10.8201 | -104.1065
10.8239 | -104.1124
10.8276 | -104.121
10.8313 | -104.1327
10.8351 | -104.1493
10.8388 | -104.1696
10.8426 | -104.198
10.8463 | -104.2292
10.8501 | -104.2617
10.8538 | -104.2997
10.8575 | -104.3367
10.8613 | -104.3759
10.865 | -104.4132
10.8688 | -104.4493
10.8725 | -104.4841
10.8763 | -104.5135
10.88 | -104.5409
10.8837 | -104.5669
10.8875 | -104.5914
10.8912 | -104.6087
10.895 | -104.6244
10.8987 | -104.6351
10.9025 | -104.6374
10.9062 | -104.6347
10.9099 | -104.6272
10.9137 | -104.6167
10.9174 | -104.6044
10.9212 | -104.5943
10.9249 | -104.5858
10.9287 | -104.5735
10.9324 | -104.5574
10.9361 | -104.5401
10.9399 | -104.5215
10.9436 | -104.5051
10.9474 | -104.4948
10.9511 | -104.4834
10.9549 | -104.4688
10.9586 | -104.4552
10.9623 | -104.4428
10.9661 | -104.4322
10.9698 | -104.422
10.9736 | -104.4122
10.9773 | -104.404
10.9811 | -104.3946
10.9848 | -104.3855
10.9885 | -104.3783
10.9923 | -104.3695
10.996 | -104.3604
10.9998 | -104.3547
11.0035 | -104.3487
11.0073 | -104.3482
11.011 | -104.3517
11.0147 | -104.3553
11.0185 | -104.3631
11.0222 | -104.3716
11.026 | -104.3841
11.0297 | -104.4032
11.0335 | -104.4271
11.0372 | -104.4526
11.0409 | -104.4784
11.0447 | -104.5086
11.0484 | -104.5392
11.0522 | -104.5722
11.0559 | -104.6058
11.0596 | -104.6396
11.0634 | -104.6712
11.0671 | -104.705
11.0709 | -104.7311
11.0746 | -104.7524
11.0784 | -104.7727
11.0821 | -104.7904
11.0858 | -104.805
11.0896 | -104.82
11.0933 | -104.8362
11.0971 | -104.846
11.1008 | -104.8475
11.1046 | -104.8464
11.1083 | -104.844
11.112 | -104.8408
11.1158 | -104.8377
11.1195 | -104.8307
11.1233 | -104.8232
11.127 | -104.8132
11.1308 | -104.8021
11.1345 | -104.7913
11.1382 | -104.7807
11.142 | -104.7746
11.1457 | -104.7711
11.1495 | -104.7684
11.1532 | -104.7668
11.157 | -104.7643
11.1607 | -104.7642
11.1644 | -104.766
11.1682 | -104.7699
11.1719 | -104.7742
11.1757 | -104.7797
11.1794 | -104.7896
11.1832 | -104.8004
11.1869 | -104.8101
11.1906 | -104.8167
11.1944 | -104.8217
11.1981 | -104.8298
11.2019 | -104.8382
11.2056 | -104.8424
11.2094 | -104.8489
11.2131 | -104.8558
11.2168 | -104.8645
11.2206 | -104.8758
11.2243 | -104.8875
11.2281 | -104.8962
11.2318 | -104.9055
11.2356 | -104.9132
11.2393 | -104.9168
11.243 | -104.9234
11.2468 | -104.9299
11.2505 | -104.9375
11.2543 | -104.9432
11.258 | -104.9508
11.2618 | -104.9601
11.2655 | -104.9674
11.2692 | -104.9675
11.273 | -104.9681
11.2767 | -104.968
11.2805 | -104.9703
11.2842 | -104.9722
11.288 | -104.9709
11.2917 | -104.9708
11.2954 | -104.97
11.2992 | -104.9624
11.3029 | -104.9493
11.3067 | -104.9345
11.3104 | -104.92
11.3142 | -104.9109
11.3179 | -104.9015
11.3216 | -104.8892
11.3254 | -104.8736
11.3291 | -104.8574
11.3329 | -104.8424
11.3366 | -104.8324
11.3404 | -104.8259
11.3441 | -104.8145
11.3478 | -104.8027
11.3516 | -104.7924
11.3553 | -104.7869
11.3591 | -104.7822
11.3628 | -104.7746
11.3666 | -104.7685
11.3703 | -104.7629
11.374 | -104.7506
11.3778 | -104.7422
11.3815 | -104.7355
11.3853 | -104.7285
11.389 | -104.7213
11.3927 | -104.7163
11.3965 | -104.7107
11.4002 | -104.7038
11.404 | -104.6959
11.4077 | -104.6851
11.4115 | -104.6752
11.4152 | -104.6624
11.4189 | -104.6524
11.4227 | -104.6417
11.4264 | -104.6262
11.4302 | -104.6136
11.4339 | -104.6034
11.4377 | -104.5883
11.4414 | -104.5718
11.4451 | -104.5526
11.4489 | -104.5344
11.4526 | -104.5197
11.4564 | -104.5066
11.4601 | -104.4962
11.4639 | -104.4908
11.4676 | -104.4859
11.4713 | -104.4843
11.4751 | -104.4869
11.4788 | -104.4877
11.4826 | -104.491
11.4863 | -104.4987
11.4901 | -104.5071
11.4938 | -104.5178
11.4975 | -104.5329
11.5013 | -104.5504
11.505 | -104.5654
11.5088 | -104.5817
11.5125 | -104.6026
11.5163 | -104.6218
11.52 | -104.639
11.5237 | -104.6564
11.5275 | -104.6734
11.5312 | -104.6919
11.535 | -104.7063
11.5387 | -104.7206
11.5425 | -104.7405
11.5462 | -104.755
11.5499 | -104.7645
11.5537 | -104.7745
11.5574 | -104.7852
11.5612 | -104.796
11.5649 | -104.8046
11.5687 | -104.8162
11.5724 | -104.8266
11.5761 | -104.8356
11.5799 | -104.8452
11.5836 | -104.8517
11.5874 | -104.8583
11.5911 | -104.8698
11.5949 | -104.8835
11.5986 | -104.8951
11.6023 | -104.9065
11.6061 | -104.9145
11.6098 | -104.9204
11.6136 | -104.9257
11.6173 | -104.9285
11.6211 | -104.9294
11.6248 | -104.9275
11.6285 | -104.9196
11.6323 | -104.9121
11.636 | -104.9023
11.6398 | -104.8887
11.6435 | -104.8704
11.6473 | -104.8543
11.651 | -104.8393
11.6547 | -104.8191
11.6585 | -104.8
11.6622 | -104.7772
11.666 | -104.7544
11.6697 | -104.7324
11.6735 | -104.7141
11.6772 | -104.6955
11.6809 | -104.6789
11.6847 | -104.6714
11.6884 | -104.6677
11.6922 | -104.665
11.6959 | -104.666
11.6996 | -104.669
11.7034 | -104.6755
11.7071 | -104.6828
11.7109 | -104.6883
11.7146 | -104.7
11.7184 | -104.7178
11.7221 | -104.7371
11.7258 | -104.7553
11.7296 | -104.7762
11.7333 | -104.799
11.7371 | -104.8215
11.7408 | -104.8452
11.7446 | -104.8665
11.7483 | -104.888
11.752 | -104.9126
11.7558 | -104.9324
11.7595 | -104.9467
11.7633 | -104.9609
11.767 | -104.9748
11.7708 | -104.9889
11.7745 | -105.0055
11.7782 | -105.0168
11.782 | -105.0214
11.7857 | -105.0191
11.7895 | -105.0133
11.7932 | -105.0048
11.797 | -104.996
11.8007 | -104.9822
11.8044 | -104.9689
11.8082 | -104.9528
11.8119 | -104.9353
11.8157 | -104.9166
11.8194 | -104.8997
11.8232 | -104.888
11.8269 | -104.8783
11.8306 | -104.8721
11.8344 | -104.863
11.8381 | -104.8537
11.8419 | -104.8464
11.8456 | -104.8467
11.8494 | -104.8484
11.8531 | -104.8469
11.8568 | -104.8461
11.8606 | -104.8485
11.8643 | -104.849
11.8681 | -104.8526
11.8718 | -104.8609
11.8756 | -104.8592
11.8793 | -104.8631
11.883 | -104.8679
11.8868 | -104.8678
11.8905 | -104.8664
11.8943 | -104.8649
11.898 | -104.8597
11.9018 | -104.8523
11.9055 | -104.8445
11.9092 | -104.8359
11.913 | -104.8267
11.9167 | -104.8151
11.9205 | -104.8028
11.9242 | -104.7885
11.928 | -104.775
11.9317 | -104.7624
11.9354 | -104.7487
11.9392 | -104.732
11.9429 | -104.7156
11.9467 | -104.703
11.9504 | -104.6904
11.9542 | -104.6784
11.9579 | -104.6704
11.9616 | -104.6628
11.9654 | -104.66
11.9691 | -104.6594
11.9729 | -104.6622
11.9766 | -104.67
11.9804 | -104.6785
11.9841 | -104.6871
11.9878 | -104.6961
11.9916 | -104.7066
11.9953 | -104.7161
11.9991 | -104.7284
12.0028 | -104.7403
12.0066 | -104.7544
12.0103 | -104.7692
12.014 | -104.7843
12.0178 | -104.7931
12.0215 | -104.8001
12.0253 | -104.8131
12.029 | -104.8242
12.0327 | -104.8329
12.0365 | -104.8449
12.0402 | -104.8552
12.044 | -104.8631
12.0477 | -104.8743
12.0515 | -104.8882
12.0552 | -104.8984
12.0589 | -104.9063
12.0627 | -104.9179
12.0664 | -104.9254
12.0702 | -104.9275
12.0739 | -104.9342
12.0777 | -104.9327
12.0814 | -104.9208
12.0851 | -104.9129
12.0889 | -104.9046
12.0926 | -104.8934
12.0964 | -104.8771
12.1001 | -104.8608
12.1039 | -104.8425
12.1076 | -104.8191
12.1113 | -104.7879
12.1151 | -104.7558
12.1188 | -104.7241
12.1226 | -104.6858
12.1263 | -104.6454
12.1301 | -104.6054
12.1338 | -104.5665
12.1375 | -104.5248
12.1413 | -104.4791
12.145 | -104.4376
12.1488 | -104.4005
12.1525 | -104.3647
12.1563 | -104.3376
12.16 | -104.3179
12.1637 | -104.3045
12.1675 | -104.2956
12.1712 | -104.2916
12.175 | -104.2917
12.1787 | -104.3001
12.1825 | -104.3168
12.1862 | -104.3406
12.1899 | -104.3696
12.1937 | -104.4064
12.1974 | -104.4494
12.2012 | -104.49
12.2049 | -104.5342
12.2087 | -104.5795
12.2124 | -104.6196
12.2161 | -104.6573
12.2199 | -104.6946
12.2236 | -104.73
12.2274 | -104.7601
12.2311 | -104.7876
12.2349 | -104.8089
12.2386 | -104.8285
12.2423 | -104.8474
12.2461 | -104.8606
12.2498 | -104.864
12.2536 | -104.8636
12.2573 | -104.864
12.2611 | -104.8527
12.2648 | -104.8363
12.2685 | -104.8195
12.2723 | -104.7993
12.276 | -104.7717
12.2798 | -104.7427
12.2835 | -104.7112
12.2873 | -104.68
12.291 | -104.6514
12.2947 | -104.6231
12.2985 | -104.597
12.3022 | -104.5695
12.306 | -104.5364
12.3097 | -104.5016
12.3134 | -104.4725
12.3172 | -104.4481
12.3209 | -104.4217
12.3247 | -104.3954
12.3284 | -104.3704
12.3322 | -104.3463
12.3359 | -104.3264
12.3396 | -104.3079
12.3434 | -104.2903
12.3471 | -104.2773
12.3509 | -104.2659
12.3546 | -104.2559
12.3584 | -104.2461
12.3621 | -104.2388
12.3658 | -104.2347
12.3696 | -104.2354
12.3733 | -104.2365
12.3771 | -104.2442
12.3808 | -104.256
12.3846 | -104.2695
12.3883 | -104.2899
12.392 | -104.3136
12.3958 | -104.3379
12.3995 | -104.3642
12.4033 | -104.3932
12.407 | -104.415
12.4108 | -104.4394
12.4145 | -104.4641
12.4182 | -104.4917
12.422 | -104.5141
12.4257 | -104.5366
12.4295 | -104.5605
12.4332 | -104.5846
12.437 | -104.6023
12.4407 | -104.6156
12.4444 | -104.6255
12.4482 | -104.6338
12.4519 | -104.6378
12.4557 | -104.6404
12.4594 | -104.6397
12.4632 | -104.637
12.4669 | -104.6325
12.4706 | -104.618
12.4744 | -104.5985
12.4781 | -104.5825
12.4819 | -104.5735
12.4856 | -104.5692
12.4894 | -104.5602
12.4931 | -104.547
12.4968 | -104.5356
12.5006 | -104.5279
12.5043 | -104.5217
12.5081 | -104.5191
12.5118 | -104.5201
12.5156 | -104.522
12.5193 | -104.526
12.523 | -104.5315
12.5268 | -104.5392
12.5305 | -104.5504
12.5343 | -104.5629
12.538 | -104.575
12.5418 | -104.5894
12.5455 | -104.6025
12.5492 | -104.6155
12.553 | -104.6312
12.5567 | -104.6446
12.5605 | -104.6587
12.5642 | -104.6714
12.568 | -104.6852
12.5717 | -104.6944
12.5754 | -104.6994
12.5792 | -104.7055
12.5829 | -104.7122
12.5867 | -104.7243
12.5904 | -104.734
12.5942 | -104.7372
12.5979 | -104.7372
12.6016 | -104.7401
12.6054 | -104.7433
12.6091 | -104.7468
12.6129 | -104.748
12.6166 | -104.7491
12.6204 | -104.752
12.6241 | -104.7547
12.6278 | -104.7599
12.6316 | -104.7646
12.6353 | -104.7708
12.6391 | -104.7791
12.6428 | -104.7834
12.6465 | -104.7883
12.6503 | -104.7994
12.654 | -104.8117
12.6578 | -104.8224
12.6615 | -104.8276
12.6653 | -104.8313
12.669 | -104.8378
12.6727 | -104.8469
12.6765 | -104.8568
12.6802 | -104.8643
12.684 | -104.8722
12.6877 | -104.8796
12.6915 | -104.8828
12.6952 | -104.8819
12.6989 | -104.881
12.7027 | -104.8812
12.7064 | -104.882
12.7102 | -104.8771
12.7139 | -104.8658
12.7177 | -104.8508
12.7214 | -104.8404
12.7251 | -104.8274
12.7289 | -104.8155
12.7326 | -104.8048
12.7364 | -104.7957
12.7401 | -104.7861
12.7439 | -104.7763
12.7476 | -104.7726
12.7513 | -104.7674
12.7551 | -104.7673
12.7588 | -104.7715
12.7626 | -104.7752
12.7663 | -104.7807
12.7701 | -104.7873
12.7738 | -104.7965
12.7775 | -104.8106
12.7813 | -104.8285
12.785 | -104.845
12.7888 | -104.8663
12.7925 | -104.888
12.7963 | -104.9106
12.8 | -104.9302
12.8037 | -104.9522
12.8075 | -104.9731
12.8112 | -104.9906
12.815 | -105.004
12.8187 | -105.0199
12.8225 | -105.0336
12.8262 | -105.0491
12.8299 | -105.0621
12.8337 | -105.0648
12.8374 | -105.0667
12.8412 | -105.0719
12.8449 | -105.078
12.8487 | -105.0816
12.8524 | -105.0793
12.8561 | -105.0758
12.8599 | -105.0742
12.8636 | -105.0739
12.8674 | -105.074
12.8711 | -105.0749
12.8749 | -105.0795
12.8786 | -105.0823
12.8823 | -105.0854
12.8861 | -105.0902
12.8898 | -105.0956
12.8936 | -105.1099
12.8973 | -105.1256
12.9011 | -105.1396
12.9048 | -105.155
12.9085 | -105.1693
12.9123 | -105.1821
12.916 | -105.1953
12.9198 | -105.2085
12.9235 | -105.2235
12.9273 | -105.2361
12.931 | -105.2457
12.9347 | -105.256
12.9385 | -105.2657
12.9422 | -105.2736
12.946 | -105.2757
12.9497 | -105.2734
12.9535 | -105.2719
12.9572 | -105.2452
12.9609 | -105.2341
12.9647 | -105.2215
12.9684 | -105.2074
12.9722 | -105.1935
12.9759 | -105.1785
12.9796 | -105.1624
12.9834 | -105.1438
12.9871 | -105.1266
12.9909 | -105.1104
12.9946 | -105.0958
12.9984 | -105.0836
13.0021 | -105.0704
13.0058 | -105.0527
13.0096 | -105.04
13.0133 | -105.0291
13.0171 | -105.0201
13.0208 | -105.013
13.0246 | -105.0076
13.0283 | -105.0027
13.032 | -104.9957
13.0358 | -104.9924
13.0395 | -104.9921
13.0433 | -104.9964
13.047 | -105.0006
13.0508 | -105.0052
13.0545 | -105.01
13.0582 | -105.0147
13.062 | -105.021
13.0657 | -105.0313
13.0695 | -105.0393
13.0732 | -105.0446
13.077 | -105.0534
13.0807 | -105.0581
13.0844 | -105.0644
13.0882 | -105.0714
13.0919 | -105.0741
13.0957 | -105.081
13.0994 | -105.0878
13.1032 | -105.0908
13.1069 | -105.0921
13.1106 | -105.0937
13.1144 | -105.0964
13.1181 | -105.0981
13.1219 | -105.095
13.1256 | -105.0921
13.1294 | -105.0901
13.1331 | -105.0872
13.1368 | -105.0838
13.1406 | -105.0804
13.1443 | -105.0783
13.1481 | -105.0757
13.1518 | -105.07
13.1556 | -105.0609
13.1593 | -105.0532
13.163 | -105.0443
13.1668 | -105.0351
13.1705 | -105.0255
13.1743 | -105.0169
13.178 | -105.0098
13.1818 | -104.9997
13.1855 | -104.9878
13.1892 | -104.9774
13.193 | -104.9717
13.1967 | -104.9682
13.2005 | -104.9587
13.2042 | -104.9492
13.208 | -104.9397
13.2117 | -104.9317
13.2154 | -104.9233
13.2192 | -104.915
13.2229 | -104.9131
13.2267 | -104.9098
13.2304 | -104.9017
13.2342 | -104.8924
13.2379 | -104.8859
13.2416 | -104.8809
13.2454 | -104.8765
13.2491 | -104.8742
13.2529 | -104.8698
13.2566 | -104.8628
13.2604 | -104.8595
13.2641 | -104.8572
13.2678 | -104.8504
13.2716 | -104.8443
13.2753 | -104.8389
13.2791 | -104.8365
13.2828 | -104.8319
13.2866 | -104.824
13.2903 | -104.8164
13.294 | -104.8075
13.2978 | -104.7969
13.3015 | -104.7871
13.3053 | -104.7764
13.309 | -104.7639
13.3127 | -104.7538
13.3165 | -104.7417
13.3202 | -104.7294
13.324 | -104.7199
13.3277 | -104.7099
13.3315 | -104.6937
13.3352 | -104.6824
13.3389 | -104.6744
13.3427 | -104.6674
13.3464 | -104.663
13.3502 | -104.66
13.3539 | -104.6608
13.3577 | -104.6663
13.3614 | -104.6754
13.3651 | -104.6832
13.3689 | -104.6887
13.3726 | -104.6985
13.3764 | -104.7107
13.3801 | -104.7252
13.3839 | -104.7407
13.3876 | -104.7584
13.3913 | -104.7784
13.3951 | -104.795
13.3988 | -104.8144
13.4026 | -104.832
13.4063 | -104.8475
13.4101 | -104.8605
13.4138 | -104.8714
13.4175 | -104.88
13.4213 | -104.8924
13.425 | -104.9031
13.4288 | -104.9094
13.4325 | -104.9135
13.4363 | -104.915
13.44 | -104.9161
13.4437 | -104.9171
13.4475 | -104.9161
13.4512 | -104.9107
13.455 | -104.9055
13.4587 | -104.8999
13.4625 | -104.8926
13.4662 | -104.8861
13.4699 | -104.8745
13.4737 | -104.8633
13.4774 | -104.8546
13.4812 | -104.8465
13.4849 | -104.8355
13.4887 | -104.8268
13.4924 | -104.8163
13.4961 | -104.8095
13.4999 | -104.7973
13.5036 | -104.7871
13.5074 | -104.7768
13.5111 | -104.7655
13.5149 | -104.7561
13.5186 | -104.7446
13.5223 | -104.7353
13.5261 | -104.729
13.5298 | -104.7195
13.5336 | -104.7083
13.5373 | -104.6966
13.5411 | -104.6848
13.5448 | -104.6721
13.5485 | -104.6597
13.5523 | -104.6448
13.556 | -104.6308
13.5598 | -104.6203
13.5635 | -104.6114
13.5673 | -104.6039
13.571 | -104.5976
13.5747 | -104.5859
13.5785 | -104.5735
13.5822 | -104.5644
13.586 | -104.5572
13.5897 | -104.5522
13.5935 | -104.5451
13.5972 | -104.5429
13.6009 | -104.5421
13.6047 | -104.5426
13.6084 | -104.5448
13.6122 | -104.5485
13.6159 | -104.5486
13.6196 | -104.5472
13.6234 | -104.5523
13.6271 | -104.5595
13.6309 | -104.568
13.6346 | -104.5774
13.6384 | -104.5863
13.6421 | -104.5969
13.6458 | -104.6099
13.6496 | -104.6229
13.6533 | -104.6363
13.6571 | -104.6509
13.6608 | -104.6639
13.6646 | -104.6807
13.6683 | -104.6959
13.672 | -104.709
13.6758 | -104.726
13.6795 | -104.7457
13.6833 | -104.7608
13.687 | -104.7772
13.6908 | -104.7934
13.6945 | -104.8081
13.6982 | -104.8268
13.702 | -104.8465
13.7057 | -104.866
13.7095 | -104.8848
13.7132 | -104.9012
13.717 | -104.9199
13.7207 | -104.9376
13.7244 | -104.9557
13.7282 | -104.9728
13.7319 | -104.9858
13.7357 | -104.9976
13.7394 | -105.0093
13.7432 | -105.0208
13.7469 | -105.029
13.7506 | -105.0377
13.7544 | -105.0429
13.7581 | -105.0443
13.7619 | -105.0432
13.7656 | -105.0375
13.7694 | -105.0298
13.7731 | -105.0263
13.7768 | -105.0209
13.7806 | -105.0097
13.7843 | -104.9998
13.7881 | -104.9904
13.7918 | -104.9811
13.7956 | -104.9695
13.7993 | -104.9584
13.803 | -104.9471
13.8068 | -104.9342
13.8105 | -104.9221
13.8143 | -104.9155
13.818 | -104.9111
13.8218 | -104.9026
13.8255 | -104.897
13.8292 | -104.8935
13.833 | -104.8864
13.8367 | -104.8788
13.8405 | -104.8745
13.8442 | -104.8682
13.848 | -104.8646
13.8517 | -104.8617
13.8554 | -104.8593
13.8592 | -104.8605
13.8629 | -104.8607
13.8667 | -104.8639
13.8704 | -104.8682
13.8742 | -104.8681
13.8779 | -104.8646
13.8816 | -104.8624
13.8854 | -104.8586
13.8891 | -104.8559
13.8929 | -104.8501
13.8966 | -104.8419
13.9004 | -104.8315
13.9041 | -104.8164
13.9078 | -104.7983
13.9116 | -104.7793
13.9153 | -104.7588
13.9191 | -104.738
13.9228 | -104.7187
13.9266 | -104.6992
13.9303 | -104.6804
13.934 | -104.6598
13.9378 | -104.6337
13.9415 | -104.6075
13.9453 | -104.5805
13.949 | -104.5536
13.9527 | -104.5292
13.9565 | -104.5054
13.9602 | -104.4817
13.964 | -104.4357
13.9677 | -104.3862
13.9715 | -104.3522
13.9752 | -104.3216
13.9789 | -104.2891
13.9827 | -104.254
13.9864 | -104.218
13.9902 | -104.1773
13.9939 | -104.136
13.9977 | -104.0947
14.0014 | -104.0537
14.0051 | -104.0128
14.0089 | -103.971
14.0126 | -103.9293
14.0164 | -103.8873
14.0201 | -103.8463
14.0239 | -103.8055
14.0276 | -103.7671
14.0313 | -103.7342
14.0351 | -103.7167
14.0388 | -103.697
14.0426 | -103.6788
14.0463 | -103.6629
14.0501 | -103.6497
14.0538 | -103.6415
14.0575 | -103.6376
14.0613 | -103.6372
14.065 | -103.6425
14.0688 | -103.6539
14.0725 | -103.6702
14.0763 | -103.6925
14.08 | -103.7223
14.0837 | -103.7539
14.0875 | -103.7876
14.0912 | -103.8228
14.095 | -103.8624
14.0987 | -103.905
14.1025 | -103.9485
14.1062 | -103.9952
14.1099 | -104.0441
14.1137 | -104.0922
14.1174 | -104.1405
14.1212 | -104.1868
14.1249 | -104.2301
14.1287 | -104.2736
14.1324 | -104.3189
14.1361 | -104.3657
14.1399 | -104.411
14.1436 | -104.4535
14.1474 | -104.4977
14.1511 | -104.5367
14.1549 | -104.5729
14.1586 | -104.6075
14.1623 | -104.6402
14.1661 | -104.6927
14.1698 | -104.741
14.1736 | -104.7634
14.1773 | -104.7782
14.1811 | -104.7922
14.1848 | -104.8045
14.1885 | -104.816
14.1923 | -104.8296
14.196 | -104.8368
14.1998 | -104.8408
14.2035 | -104.8393
14.2073 | -104.8328
14.211 | -104.8225
14.2147 | -104.811
14.2185 | -104.7984
14.2222 | -104.7836
14.226 | -104.7669
14.2297 | -104.7477
14.2334 | -104.7232
14.2372 | -104.6967
14.2409 | -104.6702
14.2447 | -104.6407
14.2484 | -104.6141
14.2522 | -104.5925
14.2559 | -104.5708
14.2596 | -104.5493
14.2634 | -104.5274
14.2671 | -104.5099
14.2709 | -104.4928
14.2746 | -104.4744
14.2784 | -104.4592
14.2821 | -104.4439
14.2858 | -104.4322
14.2896 | -104.4269
14.2933 | -104.4282
14.2971 | -104.4331
14.3008 | -104.437
14.3046 | -104.4376
14.3083 | -104.4408
14.312 | -104.4493
14.3158 | -104.4615
14.3195 | -104.4747
14.3233 | -104.4865
14.327 | -104.4981
14.3308 | -104.5111
14.3345 | -104.5244
14.3382 | -104.5359
14.342 | -104.5444
14.3457 | -104.558
14.3495 | -104.5692
14.3532 | -104.5768
14.357 | -104.5851
14.3607 | -104.5871
14.3644 | -104.5873
14.3682 | -104.5876
14.3719 | -104.586
14.3757 | -104.5809
14.3794 | -104.5759
14.3832 | -104.5719
14.3869 | -104.5685
14.3906 | -104.5665
14.3944 | -104.5635
14.3981 | -104.559
14.4019 | -104.5562
14.4056 | -104.556
14.4094 | -104.5587
14.4131 | -104.5649
14.4168 | -104.5718
14.4206 | -104.5768
14.4243 | -104.5844
14.4281 | -104.5921
14.4318 | -104.604
14.4356 | -104.6226
14.4393 | -104.6447
14.443 | -104.6707
14.4468 | -104.6972
14.4505 | -104.7265
14.4543 | -104.7556
14.458 | -104.7852
14.4618 | -104.8155
14.4655 | -104.8483
14.4692 | -104.8846
14.473 | -104.9163
14.4767 | -104.9499
14.4805 | -104.9858
14.4842 | -105.0189
14.488 | -105.0523
14.4917 | -105.0777
14.4954 | -105.098
14.4992 | -105.1121
14.5029 | -105.1237
14.5067 | -105.1381
14.5104 | -105.1535
14.5142 | -105.1626
14.5179 | -105.164
14.5216 | -105.1578
14.5254 | -105.1502
14.5291 | -105.1437
14.5329 | -105.1355
14.5366 | -105.1239
14.5404 | -105.1061
14.5441 | -105.0892
14.5478 | -105.0706
14.5516 | -105.0477
14.5553 | -105.0239
14.5591 | -105.0006
14.5628 | -104.9809
14.5665 | -104.9648
14.5703 | -104.9473
14.574 | -104.9264
14.5778 | -104.9064
14.5815 | -104.8859
14.5853 | -104.866
14.589 | -104.8498
14.5927 | -104.8343
14.5965 | -104.8181
14.6002 | -104.7998
14.604 | -104.7836
14.6077 | -104.7707
14.6115 | -104.758
14.6152 | -104.7454
14.6189 | -104.7368
14.6227 | -104.7293
14.6264 | -104.7224
14.6302 | -104.719
14.6339 | -104.7158
14.6377 | -104.7157
14.6414 | -104.7142
14.6451 | -104.7118
14.6489 | -104.7146
14.6526 | -104.7161
14.6564 | -104.7192
14.6601 | -104.7227
14.6639 | -104.73
14.6676 | -104.7404
14.6713 | -104.7552
14.6751 | -104.7708
14.6788 | -104.7849
14.6826 | -104.8037
14.6863 | -104.823
14.6901 | -104.8405
14.6938 | -104.8597
14.6975 | -104.8795
14.7013 | -104.9034
14.705 | -104.929
14.7088 | -104.9481
14.7125 | -104.9643
14.7163 | -104.9818
14.72 | -104.9998
14.7237 | -105.0168
14.7275 | -105.0326
14.7312 | -105.049
14.735 | -105.0651
14.7387 | -105.0785
14.7425 | -105.09
14.7462 | -105.1016
14.7499 | -105.1133
14.7537 | -105.1196
14.7574 | -105.1249
14.7612 | -105.1287
14.7649 | -105.1311
14.7687 | -105.1326
14.7724 | -105.131
14.7761 | -105.1292
14.7799 | -105.1273
14.7836 | -105.1249
14.7874 | -105.1266
14.7911 | -105.1277
14.7949 | -105.1263
14.7986 | -105.1242
14.8023 | -105.1218
14.8061 | -105.1216
14.8098 | -105.1217
14.8136 | -105.1228
14.8173 | -105.1247
14.8211 | -105.1232
14.8248 | -105.1263
14.8285 | -105.1294
14.8323 | -105.1294
14.836 | -105.1308
14.8398 | -105.1332
14.8435 | -105.1363
14.8473 | -105.1399
14.851 | -105.145
14.8547 | -105.1491
14.8585 | -105.1496
14.8622 | -105.1449
14.866 | -105.1421
14.8697 | -105.1385
14.8735 | -105.137
14.8772 | -105.1345
14.8809 | -105.1268
14.8847 | -105.1179
14.8884 | -105.1105
14.8922 | -105.1017
14.8959 | -105.0888
14.8996 | -105.0771
14.9034 | -105.0669
14.9071 | -105.0548
14.9109 | -105.0436
14.9146 | -105.0332
14.9184 | -105.0168
14.9221 | -105.008
14.9258 | -105.0014
14.9296 | -104.9931
14.9333 | -104.9837
14.9371 | -104.974
14.9408 | -104.966
14.9446 | -104.9595
14.9483 | -104.9515
14.952 | -104.9458
14.9558 | -104.9415
14.9595 | -104.9375
14.9633 | -104.9351
14.967 | -104.9354
14.9708 | -104.9352
14.9745 | -104.9356
14.9782 | -104.9386
14.982 | -104.9411
14.9857 | -104.9402
14.9895 | -104.9169
14.9932 | -104.8992
14.997 | -104.8976
15.0007 | -104.891
15.0044 | -104.8825
15.0082 | -104.8734
15.0119 | -104.8662
15.0157 | -104.8534
15.0194 | -104.8381
15.0232 | -104.821
15.0269 | -104.8018
15.0306 | -104.7808
15.0344 | -104.7592
15.0381 | -104.7334
15.0419 | -104.7074
15.0456 | -104.6805
15.0494 | -104.6505
15.0531 | -104.6226
15.0568 | -104.5927
15.0606 | -104.5635
15.0643 | -104.535
15.0681 | -104.51
15.0718 | -104.4902
15.0756 | -104.4732
15.0793 | -104.4557
15.083 | -104.4397
15.0868 | -104.4247
15.0905 | -104.4142
15.0943 | -104.411
15.098 | -104.4117
15.1018 | -104.4131
15.1055 | -104.4197
15.1092 | -104.4317
15.113 | -104.4448
15.1167 | -104.4723
15.1205 | -104.488
15.1242 | -104.5076
15.128 | -104.527
15.1317 | -104.5469
15.1354 | -104.5729
15.1392 | -104.5989
15.1429 | -104.6213
15.1467 | -104.642
15.1504 | -104.6637
15.1542 | -104.687
15.1579 | -104.7073
15.1616 | -104.7255
15.1654 | -104.743
15.1691 | -104.7589
15.1729 | -104.7742
15.1766 | -104.787
15.1804 | -104.7978
15.1841 | -104.8079
15.1878 | -104.819
15.1916 | -104.8274
15.1953 | -104.8342
15.1991 | -104.8386
15.2028 | -104.8458
15.2066 | -104.8713
15.2103 | -104.8836
15.214 | -104.8766
15.2178 | -104.871
15.2215 | -104.867
15.2253 | -104.8651
15.229 | -104.8616
15.2327 | -104.8603
15.2365 | -104.8587
15.2402 | -104.8578
15.244 | -104.8571
15.2477 | -104.8541
15.2515 | -104.8523
15.2552 | -104.849
15.2589 | -104.8434
15.2627 | -104.8393
15.2664 | -104.8373
15.2702 | -104.8357
15.2739 | -104.8343
15.2777 | -104.8311
15.2814 | -104.8291
15.2851 | -104.8253
15.2889 | -104.8225
15.2926 | -104.8214
15.2964 | -104.824
15.3001 | -104.8266
15.3039 | -104.8297
15.3076 | -104.8293
15.3113 | -104.8297
15.3151 | -104.8317
15.3188 | -104.8375
15.3226 | -104.8441
15.3263 | -104.85
15.3301 | -104.858
15.3338 | -104.8658
15.3375 | -104.8758
15.3413 | -104.8835
15.345 | -104.8887
15.3488 | -104.8938
15.3525 | -104.8994
15.3563 | -104.9023
15.36 | -104.9026
15.3637 | -104.9065
15.3675 | -104.9127
15.3712 | -104.9183
15.375 | -104.9222
15.3787 | -104.9226
15.3825 | -104.9203
15.3862 | -104.9176
15.3899 | -104.9157
15.3937 | -104.9111
15.3974 | -104.9028
15.4012 | -104.8894
15.4049 | -104.8753
15.4087 | -104.8581
15.4124 | -104.8413
15.4161 | -104.8235
15.4199 | -104.805
15.4236 | -104.7848
15.4274 | -104.7632
15.4311 | -104.7381
15.4349 | -104.717
15.4386 | -104.6961
15.4423 | -104.676
15.4461 | -104.6539
15.4498 | -104.6325
15.4536 | -104.6154
15.4573 | -104.5989
15.4611 | -104.5846
15.4648 | -104.5728
15.4685 | -104.5639
15.4723 | -104.5567
15.476 | -104.5525
15.4798 | -104.549
15.4835 | -104.5486
15.4873 | -104.5496
15.491 | -104.5524
15.4947 | -104.5579
15.4985 | -104.5661
15.5022 | -104.5763
15.506 | -104.5895
15.5097 | -104.5995
15.5135 | -104.6124
15.5172 | -104.6245
15.5209 | -104.6392
15.5247 | -104.6559
15.5284 | -104.6742
15.5322 | -104.6896
15.5359 | -104.7042
15.5396 | -104.7186
15.5434 | -104.7303
15.5471 | -104.7403
15.5509 | -104.7488
15.5546 | -104.754
15.5584 | -104.7579
15.5621 | -104.763
15.5658 | -104.7648
15.5696 | -104.7681
15.5733 | -104.7659
15.5771 | -104.7647
15.5808 | -104.7636
15.5846 | -104.7577
15.5883 | -104.7486
15.592 | -104.7374
15.5958 | -104.7232
15.5995 | -104.7136
15.6033 | -104.7056
15.607 | -104.6951
15.6108 | -104.68
15.6145 | -104.6654
15.6182 | -104.6545
15.622 | -104.6455
15.6257 | -104.6339
15.6295 | -104.6217
15.6332 | -104.6069
15.637 | -104.5906
15.6407 | -104.5767
15.6444 | -104.5651
15.6482 | -104.5554
15.6519 | -104.5473
15.6557 | -104.5387
15.6594 | -104.5299
15.6632 | -104.5223
15.6669 | -104.5159
15.6706 | -104.5106
15.6744 | -104.5035
15.6781 | -104.4966
15.6819 | -104.4946
15.6856 | -104.4927
15.6894 | -104.4935
15.6931 | -104.4945
15.6968 | -104.4956
15.7006 | -104.5007
15.7043 | -104.509
15.7081 | -104.5136
15.7118 | -104.5201
15.7156 | -104.529
15.7193 | -104.5381
15.723 | -104.5499
15.7268 | -104.5629
15.7305 | -104.5753
15.7343 | -104.5903
15.738 | -104.6018
15.7418 | -104.6151
15.7455 | -104.6293
15.7492 | -104.6433
15.753 | -104.66
15.7567 | -104.6766
15.7605 | -104.6924
15.7642 | -104.7074
15.768 | -104.7252
15.7717 | -104.7433
15.7754 | -104.7589
15.7792 | -104.7786
15.7829 | -104.7967
15.7867 | -104.8084
15.7904 | -104.8222
15.7942 | -104.8357
15.7979 | -104.8494
15.8016 | -104.8594
15.8054 | -104.8682
15.8091 | -104.8783
15.8129 | -104.8862
15.8166 | -104.8938
15.8204 | -104.9025
15.8241 | -104.9058
15.8278 | -104.9048
15.8316 | -104.9058
15.8353 | -104.9068
15.8391 | -104.906
15.8428 | -104.9026
15.8466 | -104.9
15.8503 | -104.8956
15.854 | -104.8915
15.8578 | -104.8861
15.8615 | -104.881
15.8653 | -104.8756
15.869 | -104.8659
15.8727 | -104.8552
15.8765 | -104.8485
15.8802 | -104.8455
15.884 | -104.8395
15.8877 | -104.8352
15.8915 | -104.8314
15.8952 | -104.8238
15.8989 | -104.8175
15.9027 | -104.8124
15.9064 | -104.805
15.9102 | -104.7955
15.9139 | -104.7883
15.9177 | -104.7804
15.9214 | -104.7733
15.9251 | -104.7665
15.9289 | -104.755
15.9326 | -104.7426
15.9364 | -104.7285
15.9401 | -104.7126
15.9439 | -104.6959
15.9476 | -104.6784
15.9513 | -104.6613
15.9551 | -104.6407
15.9588 | -104.6207
15.9626 | -104.6025
15.9663 | -104.5848
15.9701 | -104.5625
15.9738 | -104.5376
15.9775 | -104.514
15.9813 | -104.4897
15.985 | -104.4646
15.9888 | -104.4381
15.9925 | -104.4143
15.9963 | -104.393
16 | -104.3706
16.0037 | -104.3492
16.0075 | -104.329
16.0112 | -104.3104
16.015 | -104.297
16.0187 | -104.2836
16.0225 | -104.2679
16.0262 | -104.2555
16.0299 | -104.2478
16.0337 | -104.2416
16.0374 | -104.2374
16.0412 | -104.237
16.0449 | -104.2396
16.0487 | -104.2444
16.0524 | -104.2535
16.0561 | -104.2653
16.0599 | -104.2772
16.0636 | -104.2943
16.0674 | -104.3132
16.0711 | -104.3326
16.0749 | -104.3522
16.0786 | -104.3751
16.0823 | -104.3991
16.0861 | -104.4229
16.0898 | -104.4478
16.0936 | -104.4721
16.0973 | -104.4983
16.1011 | -104.524
16.1048 | -104.5487
16.1085 | -104.5739
16.1123 | -104.598
16.116 | -104.6214
16.1198 | -104.6438
16.1235 | -104.6673
16.1273 | -104.6904
16.131 | -104.7148
16.1347 | -104.7372
16.1385 | -104.7573
16.1422 | -104.7742
16.146 | -104.7861
16.1497 | -104.799
16.1534 | -104.8125
16.1572 | -104.8249
16.1609 | -104.8349
16.1647 | -104.8467
16.1684 | -104.8534
16.1722 | -104.8545
16.1759 | -104.8527
16.1796 | -104.8506
16.1834 | -104.8489
16.1871 | -104.8459
16.1909 | -104.8412
16.1946 | -104.8347
16.1984 | -104.8247
16.2021 | -104.8187
16.2058 | -104.8124
16.2096 | -104.8042
16.2133 | -104.7983
16.2171 | -104.7956
16.2208 | -104.7913
16.2246 | -104.7848
16.2283 | -104.7787
16.232 | -104.7757
16.2358 | -104.7759
16.2395 | -104.7739
16.2433 | -104.7712
16.247 | -104.7702
16.2508 | -104.7697
16.2545 | -104.772
16.2582 | -104.7731
16.262 | -104.7742
16.2657 | -104.7763
16.2695 | -104.7799
16.2732 | -104.7834
16.277 | -104.789
16.2807 | -104.796
16.2844 | -104.8032
16.2882 | -104.8099
16.2919 | -104.8154
16.2957 | -104.8217
16.2994 | -104.8313
16.3032 | -104.8424
16.3069 | -104.8546
16.3106 | -104.8648
16.3144 | -104.8755
16.3181 | -104.8883
16.3219 | -104.899
16.3256 | -104.9091
16.3294 | -104.9163
16.3331 | -104.9237
16.3368 | -104.9299
16.3406 | -104.9323
16.3443 | -104.9364
16.3481 | -104.9402
16.3518 | -104.9454
16.3556 | -104.9501
16.3593 | -104.9519
16.363 | -104.9527
16.3668 | -104.9498
16.3705 | -104.944
16.3743 | -104.9396
16.378 | -104.9365
16.3818 | -104.9299
16.3855 | -104.9205
16.3892 | -104.9106
16.393 | -104.9024
16.3967 | -104.8913
16.4005 | -104.8787
16.4042 | -104.8691
16.408 | -104.8615
16.4117 | -104.849
16.4154 | -104.8364
16.4192 | -104.8279
16.4229 | -104.8179
16.4267 | -104.8076
16.4304 | -104.7995
16.4342 | -104.7932
16.4379 | -104.7873
16.4416 | -104.7818
16.4454 | -104.7802
16.4491 | -104.7779
16.4529 | -104.7746
16.4566 | -104.7773
16.4604 | -104.7799
16.4641 | -104.7821
16.4678 | -104.7864
16.4716 | -104.7945
16.4753 | -104.8038
16.4791 | -104.814
16.4828 | -104.8219
16.4866 | -104.8307
16.4903 | -104.8404
16.494 | -104.8504
16.4978 | -104.8598
16.5015 | -104.8692
16.5053 | -104.8809
16.509 | -104.8927
16.5127 | -104.9025
16.5165 | -104.9099
16.5202 | -104.9176
16.524 | -104.9271
16.5277 | -104.9371
16.5315 | -104.9451
16.5352 | -104.9527
16.5389 | -104.9626
16.5427 | -104.9704
16.5464 | -104.978
16.5502 | -104.985
16.5539 | -104.994
16.5577 | -105.0051
16.5614 | -105.0144
16.5651 | -105.0246
16.5689 | -105.0327
16.5726 | -105.0402
16.5764 | -105.0486
16.5801 | -105.0554
16.5839 | -105.0631
16.5876 | -105.0706
16.5913 | -105.0766
16.5951 | -105.0813
16.5988 | -105.0862
16.6026 | -105.0922
16.6063 | -105.0991
16.6101 | -105.1029
16.6138 | -105.1057
16.6175 | -105.1091
16.6213 | -105.1105
16.625 | -105.1141
16.6288 | -105.1148
16.6325 | -105.1155
16.6363 | -105.1167
16.64 | -105.1139
16.6437 | -105.1095
16.6475 | -105.103
16.6512 | -105.095
16.655 | -105.0866
16.6587 | -105.0772
16.6625 | -105.0715
16.6662 | -105.0649
16.6699 | -105.0543
16.6737 | -105.0421
16.6774 | -105.032
16.6812 | -105.0192
16.6849 | -105.0041
16.6887 | -104.9918
16.6924 | -104.9805
16.6961 | -104.9702
16.6999 | -104.9607
16.7036 | -104.9506
16.7074 | -104.9393
16.7111 | -104.9317
16.7149 | -104.9278
16.7186 | -104.9252
16.7223 | -104.9224
16.7261 | -104.9178
16.7298 | -104.9159
16.7336 | -104.9156
16.7373 | -104.923
16.7411 | -104.9253
16.7448 | -104.9262
16.7485 | -104.929
16.7523 | -104.9339
16.756 | -104.9424
16.7598 | -104.9499
16.7635 | -104.9585
16.7673 | -104.9669
16.771 | -104.9758
16.7747 | -104.9832
16.7785 | -104.99
16.7822 | -104.9996
16.786 | -105.0097
16.7897 | -105.0219
16.7934 | -105.032
16.7972 | -105.0396
16.8009 | -105.0481
16.8047 | -105.0586
16.8084 | -105.0695
16.8122 | -105.0782
16.8159 | -105.0858
16.8196 | -105.0913
16.8234 | -105.0959
16.8271 | -105.1015
16.8309 | -105.1081
16.8346 | -105.1153
16.8384 | -105.1215
16.8421 | -105.1269
16.8458 | -105.1335
16.8496 | -105.1388
16.8533 | -105.1417
16.8571 | -105.1437
16.8608 | -105.1478
16.8646 | -105.1482
16.8683 | -105.1492
16.872 | -105.1499
16.8758 | -105.1512
16.8795 | -105.1519
16.8833 | -105.153
16.887 | -105.1549
16.8908 | -105.152
16.8945 | -105.1435
16.8982 | -105.1374
16.902 | -105.1303
16.9057 | -105.1198
16.9095 | -105.1088
16.9132 | -105.0972
16.917 | -105.0829
16.9207 | -105.0708
16.9244 | -105.0593
16.9282 | -105.0436
16.9319 | -105.0283
16.9357 | -105.0143
16.9394 | -104.9998
16.9432 | -104.9834
16.9469 | -104.9655
16.9506 | -104.9493
16.9544 | -104.934
16.9581 | -104.9189
16.9619 | -104.9048
16.9656 | -104.893
16.9694 | -104.8806
16.9731 | -104.867
16.9768 | -104.8541
16.9806 | -104.8443
16.9843 | -104.835
16.9881 | -104.8262
16.9918 | -104.8182
16.9956 | -104.8116
16.9993 | -104.8045
17.003 | -104.7991
17.0068 | -104.7956
17.0105 | -104.7927
17.0143 | -104.7868
17.018 | -104.7832
17.0218 | -104.783
17.0255 | -104.7831
17.0292 | -104.7826
17.033 | -104.7841
17.0367 | -104.7869
17.0405 | -104.79
17.0442 | -104.7965
17.048 | -104.8035
17.0517 | -104.8102
17.0554 | -104.8181
17.0592 | -104.8273
17.0629 | -104.8372
17.0667 | -104.8453
17.0704 | -104.8521
17.0742 | -104.8586
17.0779 | -104.8661
17.0816 | -104.8707
17.0854 | -104.8755
17.0891 | -104.8784
17.0929 | -104.8797
17.0966 | -104.8837
17.1004 | -104.8857
17.1041 | -104.8853
17.1078 | -104.8856
17.1116 | -104.8832
17.1153 | -104.8799
17.1191 | -104.8776
17.1228 | -104.8754
17.1266 | -104.8697
17.1303 | -104.8633
17.134 | -104.8583
17.1378 | -104.8528
17.1415 | -104.8473
17.1453 | -104.8421
17.149 | -104.8362
17.1527 | -104.8303
17.1565 | -104.8248
17.1602 | -104.8159
17.164 | -104.805
17.1677 | -104.795
17.1715 | -104.7874
17.1752 | -104.7744
17.1789 | -104.7597
17.1827 | -104.7448
17.1864 | -104.7313
17.1902 | -104.7178
17.1939 | -104.7014
17.1977 | -104.6858
17.2014 | -104.6719
17.2051 | -104.6574
17.2089 | -104.6414
17.2126 | -104.6253
17.2164 | -104.611
17.2201 | -104.5972
17.2239 | -104.5832
17.2276 | -104.5671
17.2313 | -104.5532
17.2351 | -104.5428
17.2388 | -104.5337
17.2426 | -104.5257
17.2463 | -104.5206
17.2501 | -104.5167
17.2538 | -104.514
17.2575 | -104.5161
17.2613 | -104.5198
17.265 | -104.5218
17.2688 | -104.5263
17.2725 | -104.5345
17.2763 | -104.5513
17.28 | -104.5584
17.2837 | -104.5683
17.2875 | -104.5808
17.2912 | -104.5906
17.295 | -104.6013
17.2987 | -104.6138
17.3025 | -104.6282
17.3062 | -104.6421
17.3099 | -104.6547
17.3137 | -104.6686
17.3174 | -104.682
17.3212 | -104.6971
17.3249 | -104.7092
17.3287 | -104.7211
17.3324 | -104.7341
17.3361 | -104.7462
17.3399 | -104.7574
17.3436 | -104.7659
17.3474 | -104.7739
17.3511 | -104.78
17.3549 | -104.7821
17.3586 | -104.7851
17.3623 | -104.789
17.3661 | -104.7901
17.3698 | -104.7883
17.3736 | -104.7874
17.3773 | -104.7845
17.3811 | -104.7818
17.3848 | -104.7767
17.3885 | -104.7673
17.3923 | -104.7563
17.396 | -104.7485
17.3998 | -104.7411
17.4035 | -104.7315
17.4073 | -104.7255
17.411 | -104.7173
17.4147 | -104.7086
17.4185 | -104.7005
17.4222 | -104.6929
17.426 | -104.6865
17.4297 | -104.6827
17.4335 | -104.681
17.4372 | -104.6778
17.4409 | -104.6744
17.4447 | -104.6715
17.4484 | -104.6699
17.4522 | -104.67
17.4559 | -104.6702
17.4596 | -104.6729
17.4634 | -104.6768
17.4671 | -104.682
17.4709 | -104.6885
17.4746 | -104.6978
17.4784 | -104.7083
17.4821 | -104.7174
17.4858 | -104.7264
17.4896 | -104.7378
17.4933 | -104.7512
17.4971 | -104.763
17.5008 | -104.7786
17.5046 | -104.7946
17.5083 | -104.8059
17.512 | -104.8155
17.5158 | -104.8255
17.5195 | -104.8362
17.5233 | -104.8478
17.527 | -104.8606
17.5308 | -104.8726
17.5345 | -104.8843
17.5382 | -104.8948
17.542 | -104.906
17.5457 | -104.9189
17.5495 | -104.9298
17.5532 | -104.9398
17.557 | -104.9496
17.5607 | -104.9614
17.5644 | -104.9714
17.5682 | -104.9819
17.5719 | -104.9908
17.5757 | -104.9995
17.5794 | -105.0087
17.5832 | -105.0137
17.5869 | -105.0151
17.5906 | -105.0168
17.5944 | -105.02
17.5981 | -105.024
17.6019 | -105.0263
17.6056 | -105.0301
17.6094 | -105.0327
17.6131 | -105.0299
17.6168 | -105.0255
17.6206 | -105.0222
17.6243 | -105.0189
17.6281 | -105.0191
17.6318 | -105.0168
17.6356 | -105.0148
17.6393 | -105.0129
17.643 | -105.0071
17.6468 | -105.001
17.6505 | -104.9969
17.6543 | -104.9946
17.658 | -104.9903
17.6618 | -104.984
17.6655 | -104.9779
17.6692 | -104.9713
17.673 | -104.967
17.6767 | -104.9632
17.6805 | -104.9587
17.6842 | -104.9565
17.688 | -104.9547
17.6917 | -104.9517
17.6954 | -104.9501
17.6992 | -104.9482
17.7029 | -104.9399
17.7067 | -104.9343
17.7104 | -104.9268
17.7142 | -104.9195
17.7179 | -104.912
17.7216 | -104.9042
17.7254 | -104.8926
17.7291 | -104.8777
17.7329 | -104.866
17.7366 | -104.8577
17.7404 | -104.8464
17.7441 | -104.8339
17.7478 | -104.8232
17.7516 | -104.809
17.7553 | -104.7936
17.7591 | -104.781
17.7628 | -104.7715
17.7665 | -104.7574
17.7703 | -104.7389
17.774 | -104.7212
17.7778 | -104.7058
17.7815 | -104.6897
17.7853 | -104.67
17.789 | -104.6528
17.7927 | -104.6364
17.7965 | -104.6143
17.8002 | -104.595
17.804 | -104.5795
17.8077 | -104.5643
17.8115 | -104.5485
17.8152 | -104.5397
17.8189 | -104.528
17.8227 | -104.513
17.8264 | -104.5006
17.8302 | -104.4899
17.8339 | -104.4793
17.8377 | -104.4725
17.8414 | -104.4707
17.8451 | -104.4705
17.8489 | -104.468
17.8526 | -104.4666
17.8564 | -104.4679
17.8601 | -104.4686
17.8639 | -104.4692
17.8676 | -104.472
17.8713 | -104.4743
17.8751 | -104.4779
17.8788 | -104.482
17.8826 | -104.4851
17.8863 | -104.4908
17.8901 | -104.4983
17.8938 | -104.5055
17.8975 | -104.514
17.9013 | -104.5227
17.905 | -104.5303
17.9088 | -104.5371
17.9125 | -104.5449
17.9163 | -104.5554
17.92 | -104.5639
17.9237 | -104.5729
17.9275 | -104.5807
17.9312 | -104.5871
17.935 | -104.5913
17.9387 | -104.5927
17.9425 | -104.598
17.9462 | -104.6062
17.9499 | -104.6116
17.9537 | -104.6174
17.9574 | -104.626
17.9612 | -104.6324
17.9649 | -104.6362
17.9687 | -104.6368
17.9724 | -104.6364
17.9761 | -104.6345
17.9799 | -104.6372
17.9836 | -104.6408
17.9874 | -104.6394
17.9911 | -104.6342
17.9949 | -104.627
17.9986 | -104.617
18.0023 | -104.6086
18.0061 | -104.6017
18.0098 | -104.5958
18.0136 | -104.5879
18.0173 | -104.5769
18.0211 | -104.5693
18.0248 | -104.5639
18.0285 | -104.5549
18.0323 | -104.5411
18.036 | -104.5287
18.0398 | -104.5201
18.0435 | -104.5089
18.0473 | -104.4966
18.051 | -104.4897
18.0547 | -104.4829
18.0585 | -104.4725
18.0622 | -104.4633
18.066 | -104.4582
18.0697 | -104.4552
18.0735 | -104.4516
18.0772 | -104.4475
18.0809 | -104.4478
18.0847 | -104.4488
18.0884 | -104.4516
18.0922 | -104.4526
18.0959 | -104.451
18.0996 | -104.452
18.1034 | -104.4563
18.1071 | -104.463
18.1109 | -104.4695
18.1146 | -104.4779
18.1184 | -104.4871
18.1221 | -104.4937
18.1258 | -104.4993
18.1296 | -104.5062
18.1333 | -104.5112
18.1371 | -104.5139
18.1408 | -104.5193
18.1446 | -104.5243
18.1483 | -104.5258
18.152 | -104.5262
18.1558 | -104.5262
18.1595 | -104.5235
18.1633 | -104.5186
18.167 | -104.513
18.1708 | -104.5052
18.1745 | -104.4966
18.1782 | -104.4851
18.182 | -104.4725
18.1857 | -104.4588
18.1895 | -104.4425
18.1932 | -104.4271
18.197 | -104.4118
18.2007 | -104.3923
18.2044 | -104.3696
18.2082 | -104.3473
18.2119 | -104.3249
18.2157 | -104.3035
18.2194 | -104.2831
18.2232 | -104.2636
18.2269 | -104.2449
18.2306 | -104.2259
18.2344 | -104.2056
18.2381 | -104.1851
18.2419 | -104.1659
18.2456 | -104.1471
18.2494 | -104.1322
18.2531 | -104.1232
18.2568 | -104.1135
18.2606 | -104.1043
18.2643 | -104.0967
18.2681 | -104.09
18.2718 | -104.0868
18.2756 | -104.0844
18.2793 | -104.0833
18.283 | -104.0835
18.2868 | -104.0855
18.2905 | -104.0928
18.2943 | -104.1024
18.298 | -104.11
18.3018 | -104.1148
18.3055 | -104.1191
18.3092 | -104.1248
18.313 | -104.1324
18.3167 | -104.1412
18.3205 | -104.1503
18.3242 | -104.1555
18.328 | -104.1609
18.3317 | -104.1671
18.3354 | -104.1749
18.3392 | -104.1831
18.3429 | -104.1887
18.3467 | -104.1932
18.3504 | -104.1965
18.3542 | -104.2009
18.3579 | -104.1974
18.3616 | -104.1986
18.3654 | -104.1972
18.3691 | -104.195
18.3729 | -104.1927
18.3766 | -104.1906
18.3804 | -104.1873
18.3841 | -104.1878
18.3878 | -104.1904
18.3916 | -104.1903
18.3953 | -104.1891
18.3991 | -104.1896
18.4028 | -104.1857
18.4065 | -104.1805
18.4103 | -104.1792
18.414 | -104.1778
18.4178 | -104.1765
18.4215 | -104.1764
18.4253 | -104.1755
18.429 | -104.1743
18.4327 | -104.1751
18.4365 | -104.1761
18.4402 | -104.1794
18.444 | -104.1832
18.4477 | -104.1879
18.4515 | -104.1934
18.4552 | -104.1988
18.4589 | -104.2049
18.4627 | -104.2114
18.4664 | -104.2201
18.4702 | -104.2293
18.4739 | -104.2357
18.4777 | -104.2405
18.4814 | -104.2433
18.4851 | -104.2499
18.4889 | -104.2585
18.4926 | -104.266
18.4964 | -104.2729
18.5001 | -104.2778
18.5039 | -104.2829
18.5076 | -104.2899
18.5113 | -104.2943
18.5151 | -104.2948
18.5188 | -104.2973
18.5226 | -104.2981
18.5263 | -104.2953
18.5301 | -104.2887
18.5338 | -104.2828
18.5375 | -104.2785
18.5413 | -104.2713
18.545 | -104.2619
18.5488 | -104.2523
18.5525 | -104.24
18.5563 | -104.2285
18.56 | -104.2195
18.5637 | -104.214
18.5675 | -104.2064
18.5712 | -104.1943
18.575 | -104.1809
18.5787 | -104.169
18.5825 | -104.1594
18.5862 | -104.1508
18.5899 | -104.1412
18.5937 | -104.132
18.5974 | -104.1204
18.6012 | -104.1061
18.6049 | -104.0947
18.6087 | -104.0861
18.6124 | -104.0804
18.6161 | -104.0741
18.6199 | -104.0649
18.6236 | -104.0543
18.6274 | -104.048
18.6311 | -104.043
18.6349 | -104.0382
18.6386 | -104.0328
18.6423 | -104.028
18.6461 | -104.0243
18.6498 | -104.0187
18.6536 | -104.0144
18.6573 | -104.0122
18.6611 | -104.0069
18.6648 | -103.9999
18.6685 | -103.997
18.6723 | -103.9949
18.676 | -103.9907
18.6798 | -103.9887
18.6835 | -103.9881
18.6873 | -103.9857
18.691 | -103.9843
18.6947 | -103.9823
18.6985 | -103.9789
18.7022 | -103.9745
18.706 | -103.9716
18.7097 | -103.9672
18.7135 | -103.9606
18.7172 | -103.9531
18.7209 | -103.9441
18.7247 | -103.9365
18.7284 | -103.9317
18.7322 | -103.9269
18.7359 | -103.9189
18.7396 | -103.9129
18.7434 | -103.908
18.7471 | -103.9035
18.7509 | -103.8984
18.7546 | -103.8911
18.7584 | -103.8848
18.7621 | -103.8816
18.7658 | -103.8774
18.7696 | -103.8686
18.7733 | -103.8588
18.7771 | -103.8512
18.7808 | -103.8441
18.7846 | -103.8381
18.7883 | -103.8319
18.792 | -103.8234
18.7958 | -103.8167
18.7995 | -103.8086
18.8033 | -103.7997
18.807 | -103.791
18.8108 | -103.7842
18.8145 | -103.7754
18.8182 | -103.7647
18.822 | -103.7507
18.8257 | -103.7379
18.8295 | -103.7281
18.8332 | -103.7211
18.837 | -103.7138
18.8407 | -103.7047
18.8444 | -103.6931
18.8482 | -103.6827
18.8519 | -103.6738
18.8557 | -103.6653
18.8594 | -103.6587
18.8632 | -103.6563
18.8669 | -103.6566
18.8706 | -103.6571
18.8744 | -103.6567
18.8781 | -103.6574
18.8819 | -103.66
18.8856 | -103.6662
18.8894 | -103.6745
18.8931 | -103.6818
18.8968 | -103.6982
18.9006 | -103.7017
18.9043 | -103.7107
18.9081 | -103.7228
18.9118 | -103.7343
18.9156 | -103.748
18.9193 | -103.7614
18.923 | -103.7716
18.9268 | -103.7794
18.9305 | -103.7915
18.9343 | -103.8072
18.938 | -103.8206
18.9418 | -103.8346
18.9455 | -103.8491
18.9492 | -103.8624
18.953 | -103.8713
18.9567 | -103.8747
18.9605 | -103.8826
18.9642 | -103.8923
18.968 | -103.9007
18.9717 | -103.908
18.9754 | -103.9107
18.9792 | -103.9106
18.9829 | -103.9113
18.9867 | -103.91
18.9904 | -103.9087
18.9942 | -103.9043
18.9979 | -103.8959
19.0016 | -103.8871
19.0054 | -103.8794
19.0091 | -103.8704
19.0129 | -103.8583
19.0166 | -103.8438
19.0204 | -103.8253
19.0241 | -103.8066
19.0278 | -103.7869
19.0316 | -103.7666
19.0353 | -103.7485
19.0391 | -103.7328
19.0428 | -103.7164
19.0465 | -103.6965
19.0503 | -103.6784
19.054 | -103.661
19.0578 | -103.6439
19.0615 | -103.6305
19.0653 | -103.618
19.069 | -103.6065
19.0727 | -103.5959
19.0765 | -103.5877
19.0802 | -103.5762
19.084 | -103.5703
19.0877 | -103.5689
19.0915 | -103.5702
19.0952 | -103.5711
19.0989 | -103.5716
19.1027 | -103.5724
19.1064 | -103.5782
19.1102 | -103.5875
19.1139 | -103.6006
19.1177 | -103.6125
19.1214 | -103.6257
19.1251 | -103.6421
19.1289 | -103.6574
19.1326 | -103.6717
19.1364 | -103.6861
19.1401 | -103.7005
19.1439 | -103.7167
19.1476 | -103.7326
19.1513 | -103.7477
19.1551 | -103.7625
19.1588 | -103.7781
19.1626 | -103.7924
19.1663 | -103.8068
19.1701 | -103.8246
19.1738 | -103.8427
19.1775 | -103.8585
19.1813 | -103.8714
19.185 | -103.8819
19.1888 | -103.8918
19.1925 | -103.9029
19.1963 | -103.9129
19.2 | -103.9251
19.2037 | -103.9363
19.2075 | -103.943
19.2112 | -103.9479
19.215 | -103.9528
19.2187 | -103.9575
19.2225 | -103.9632
19.2262 | -103.973
19.2299 | -103.9878
19.2337 | -103.995
19.2374 | -103.9995
19.2412 | -104.0029
19.2449 | -104.0021
19.2487 | -104.0017
19.2524 | -104.0059
19.2561 | -104.0109
19.2599 | -104.0146
19.2636 | -104.0136
19.2674 | -104.0085
19.2711 | -104.0026
19.2749 | -103.9942
19.2786 | -103.986
19.2823 | -103.9784
19.2861 | -103.9711
19.2898 | -103.9607
19.2936 | -103.9486
19.2973 | -103.9371
19.3011 | -103.9264
19.3048 | -103.9142
19.3085 | -103.8998
19.3123 | -103.8835
19.316 | -103.8658
19.3198 | -103.8487
19.3235 | -103.8296
19.3273 | -103.8106
19.331 | -103.7931
19.3347 | -103.7742
19.3385 | -103.7534
19.3422 | -103.7317
19.346 | -103.7112
19.3497 | -103.692
19.3534 | -103.6766
19.3572 | -103.6594
19.3609 | -103.6418
19.3647 | -103.6255
19.3684 | -103.6102
19.3722 | -103.5978
19.3759 | -103.5878
19.3796 | -103.5773
19.3834 | -103.5696
19.3871 | -103.5639
19.3909 | -103.5611
19.3946 | -103.5623
19.3984 | -103.5621
19.4021 | -103.5648
19.4058 | -103.5735
19.4096 | -103.5848
19.4133 | -103.5976
19.4171 | -103.6109
19.4208 | -103.6269
19.4246 | -103.6456
19.4283 | -103.6694
19.432 | -103.6938
19.4358 | -103.7327
19.4395 | -103.7594
19.4433 | -103.7876
19.447 | -103.8175
19.4508 | -103.8495
19.4545 | -103.8796
19.4582 | -103.9124
19.462 | -103.9495
19.4657 | -103.986
19.4695 | -104.0205
19.4732 | -104.0544
19.477 | -104.0858
19.4807 | -104.1161
19.4844 | -104.1486
19.4882 | -104.1787
19.4919 | -104.2059
19.4957 | -104.2319
19.4994 | -104.2566
19.5032 | -104.2791
19.5069 | -104.2991
19.5106 | -104.3184
19.5144 | -104.3354
19.5181 | -104.3524
19.5219 | -104.3722
19.5256 | -104.3894
19.5294 | -104.3976
19.5331 | -104.4026
19.5368 | -104.4085
19.5406 | -104.4136
19.5443 | -104.4198
19.5481 | -104.4226
19.5518 | -104.4223
19.5556 | -104.4206
19.5593 | -104.4155
19.563 | -104.4104
19.5668 | -104.4067
19.5705 | -104.4022
19.5743 | -104.395
19.578 | -104.385
19.5818 | -104.3715
19.5855 | -104.3596
19.5892 | -104.3502
19.593 | -104.3431
19.5967 | -104.3356
19.6005 | -104.3285
19.6042 | -104.3173
19.608 | -104.305
19.6117 | -104.2962
19.6154 | -104.2892
19.6192 | -104.2807
19.6229 | -104.2705
19.6267 | -104.2592
19.6304 | -104.2488
19.6342 | -104.2342
19.6379 | -104.2189
19.6416 | -104.2032
19.6454 | -104.1869
19.6491 | -104.1699
19.6529 | -104.1538
19.6566 | -104.1371
19.6604 | -104.1178
19.6641 | -104.0986
19.6678 | -104.0788
19.6716 | -104.0597
19.6753 | -104.0433
19.6791 | -104.0253
19.6828 | -104.004
19.6866 | -103.9831
19.6903 | -103.9629
19.694 | -103.9433
19.6978 | -103.9273
19.7015 | -103.9114
19.7053 | -103.893
19.709 | -103.877
19.7127 | -103.8624
19.7165 | -103.8465
19.7202 | -103.8326
19.724 | -103.8202
19.7277 | -103.8079
19.7315 | -103.7955
19.7352 | -103.7848
19.7389 | -103.7774
19.7427 | -103.7695
19.7464 | -103.7625
19.7502 | -103.7577
19.7539 | -103.7547
19.7577 | -103.7528
19.7614 | -103.7524
19.7651 | -103.7518
19.7689 | -103.7536
19.7726 | -103.7547
19.7764 | -103.756
19.7801 | -103.7573
19.7839 | -103.7595
19.7876 | -103.7616
19.7913 | -103.7646
19.7951 | -103.7682
19.7988 | -103.7729
19.8026 | -103.7778
19.8063 | -103.783
19.8101 | -103.7899
19.8138 | -103.7947
19.8175 | -103.7987
19.8213 | -103.8038
19.825 | -103.8097
19.8288 | -103.8158
19.8325 | -103.8218
19.8363 | -103.8252
19.84 | -103.8275
19.8437 | -103.8292
19.8475 | -103.8292
19.8512 | -103.8278
19.855 | -103.8255
19.8587 | -103.823
19.8625 | -103.8209
19.8662 | -103.8166
19.8699 | -103.8083
19.8737 | -103.7963
19.8774 | -103.7822
19.8812 | -103.7694
19.8849 | -103.7569
19.8887 | -103.7451
19.8924 | -103.7319
19.8961 | -103.7179
19.8999 | -103.705
19.9036 | -103.6906
19.9074 | -103.6751
19.9111 | -103.6597
19.9149 | -103.6462
19.9186 | -103.633
19.9223 | -103.6195
19.9261 | -103.6057
19.9298 | -103.5909
19.9336 | -103.5726
19.9373 | -103.5572
19.9411 | -103.5426
19.9448 | -103.5275
19.9485 | -103.5145
19.9523 | -103.5007
19.956 | -103.486
19.9598 | -103.4704
19.9635 | -103.4567
19.9673 | -103.4447
19.971 | -103.4347
19.9747 | -103.4279
19.9785 | -103.4184
19.9822 | -103.4087
19.986 | -103.4004
19.9897 | -103.3934
19.9934 | -103.3873
19.9972 | -103.383
20.0009 | -103.3797
20.0047 | -103.3758
20.0084 | -103.373
20.0122 | -103.3733
20.0159 | -103.3751
20.0196 | -103.3786
20.0234 | -103.3796
20.0271 | -103.3817
20.0309 | -103.3865
20.0346 | -103.3916
20.0384 | -103.3973
20.0421 | -103.4027
20.0458 | -103.4057
20.0496 | -103.4093
20.0533 | -103.4163
20.0571 | -103.4263
20.0608 | -103.4343
20.0646 | -103.4391
20.0683 | -103.4452
20.072 | -103.4538
20.0758 | -103.4601
20.0795 | -103.4672
20.0833 | -103.475
20.087 | -103.482
20.0908 | -103.4888
20.0945 | -103.4954
20.0982 | -103.5018
20.102 | -103.5086
20.1057 | -103.5151
20.1095 | -103.5199
20.1132 | -103.5263
20.117 | -103.5311
20.1207 | -103.5366
20.1244 | -103.5421
20.1282 | -103.5463
20.1319 | -103.5514
20.1357 | -103.5562
20.1394 | -103.5605
20.1432 | -103.566
20.1469 | -103.5703
20.1506 | -103.5733
20.1544 | -103.5778
20.1581 | -103.5853
20.1619 | -103.5921
20.1656 | -103.5974
20.1694 | -103.6035
20.1731 | -103.6075
20.1768 | -103.6108
20.1806 | -103.6139
20.1843 | -103.6183
20.1881 | -103.623
20.1918 | -103.6288
20.1956 | -103.6352
20.1993 | -103.6448
20.203 | -103.6534
20.2068 | -103.663
20.2105 | -103.6738
20.2143 | -103.6857
20.218 | -103.701
20.2218 | -103.7151
20.2255 | -103.7299
20.2292 | -103.7462
20.233 | -103.7607
20.2367 | -103.7754
20.2405 | -103.7945
20.2442 | -103.8165
20.248 | -103.8334
20.2517 | -103.8488
20.2554 | -103.8676
20.2592 | -103.8864
20.2629 | -103.9053
20.2667 | -103.9232
20.2704 | -103.94
20.2742 | -103.9549
20.2779 | -103.969
20.2816 | -103.9839
20.2854 | -103.9976
20.2891 | -104.008
20.2929 | -104.0209
20.2966 | -104.0348
20.3004 | -104.0451
20.3041 | -104.0522
20.3078 | -104.0574
20.3116 | -104.0661
20.3153 | -104.074
20.3191 | -104.0797
20.3228 | -104.0843
20.3266 | -104.0879
20.3303 | -104.0908
20.334 | -104.0944
20.3378 | -104.0955
20.3415 | -104.092
20.3453 | -104.0821
20.349 | -104.075
20.3527 | -104.07
20.3565 | -104.0623
20.3602 | -104.0517
20.364 | -104.0399
20.3677 | -104.0249
20.3715 | -104.0087
20.3752 | -103.9923
20.3789 | -103.9761
20.3827 | -103.956
20.3864 | -103.9348
20.3902 | -103.9145
20.3939 | -103.894
20.3977 | -103.8719
20.4014 | -103.8478
20.4051 | -103.8244
20.4089 | -103.8
20.4126 | -103.7749
20.4164 | -103.7474
20.4201 | -103.7203
20.4239 | -103.6935
20.4276 | -103.6668
20.4313 | -103.6389
20.4351 | -103.6116
20.4388 | -103.5851
20.4426 | -103.5595
20.4463 | -103.5332
20.4501 | -103.5071
20.4538 | -103.4803
20.4575 | -103.454
20.4613 | -103.4318
20.465 | -103.4133
20.4688 | -103.3941
20.4725 | -103.3724
20.4763 | -103.3536
20.48 | -103.3348
20.4837 | -103.318
20.4875 | -103.2999
20.4912 | -103.2841
20.495 | -103.2684
20.4987 | -103.2513
20.5025 | -103.237
20.5062 | -103.2256
20.5099 | -103.2148
20.5137 | -103.2048
20.5174 | -103.2024
20.5212 | -103.1937
20.5249 | -103.1869
20.5287 | -103.1791
20.5324 | -103.1712
20.5361 | -103.1654
20.5399 | -103.1629
20.5436 | -103.1589
20.5474 | -103.1515
20.5511 | -103.1449
20.5549 | -103.1373
20.5586 | -103.1338
20.5623 | -103.13
20.5661 | -103.1263
20.5698 | -103.124
20.5736 | -103.1195
20.5773 | -103.1149
20.5811 | -103.1128
20.5848 | -103.1098
20.5885 | -103.1041
20.5923 | -103.0994
20.596 | -103.0961
20.5998 | -103.0913
20.6035 | -103.0831
20.6073 | -103.0764
20.611 | -103.0676
20.6147 | -103.0593
20.6185 | -103.0526
20.6222 | -103.0464
20.626 | -103.0406
20.6297 | -103.0355
20.6334 | -103.0314
20.6372 | -103.0288
20.6409 | -103.0252
20.6447 | -103.02
20.6484 | -103.0134
20.6522 | -103.009
20.6559 | -103.0053
20.6596 | -103.0016
20.6634 | -102.9962
20.6671 | -102.9924
20.6709 | -102.9882
20.6746 | -102.9866
20.6784 | -102.9837
20.6821 | -102.9792
20.6858 | -102.9751
20.6896 | -102.9734
20.6933 | -102.972
20.6971 | -102.9713
20.7008 | -102.9707
20.7046 | -102.9711
20.7083 | -102.9743
20.712 | -102.9763
20.7158 | -102.9764
20.7195 | -102.9775
20.7233 | -102.9802
20.727 | -102.9834
20.7308 | -102.9875
20.7345 | -102.9903
20.7382 | -102.9923
20.742 | -102.9996
20.7457 | -103.0078
20.7495 | -103.0168
20.7532 | -103.024
20.757 | -103.0302
20.7607 | -103.0377
20.7644 | -103.0456
20.7682 | -103.0537
20.7719 | -103.063
20.7757 | -103.0724
20.7794 | -103.0822
20.7832 | -103.0892
20.7869 | -103.0968
20.7906 | -103.1084
20.7944 | -103.1159
20.7981 | -103.1199
20.8019 | -103.1227
20.8056 | -103.1274
20.8094 | -103.1304
20.8131 | -103.1338
20.8168 | -103.1391
20.8206 | -103.1418
20.8243 | -103.1446
20.8281 | -103.1474
20.8318 | -103.1484
20.8356 | -103.1473
20.8393 | -103.1466
20.843 | -103.1461
20.8468 | -103.1447
20.8505 | -103.144
20.8543 | -103.1374
20.858 | -103.1291
20.8618 | -103.1223
20.8655 | -103.1149
20.8692 | -103.1081
20.873 | -103.1012
20.8767 | -103.0939
20.8805 | -103.0859
20.8842 | -103.0777
20.888 | -103.0692
20.8917 | -103.0614
20.8954 | -103.0542
20.8992 | -103.048
20.9029 | -103.0395
20.9067 | -103.0334
20.9104 | -103.0281
20.9142 | -103.0211
20.9179 | -103.0153
20.9216 | -103.0108
20.9254 | -103.0074
20.9291 | -103.0049
20.9329 | -103.0038
20.9366 | -103.0029
20.9404 | -103.0028
20.9441 | -103.0055
20.9478 | -103.0064
20.9516 | -103.0088
20.9553 | -103.0143
20.9591 | -103.0225
20.9628 | -103.029
20.9666 | -103.0371
20.9703 | -103.0454
20.974 | -103.0574
20.9778 | -103.0679
20.9815 | -103.0786
20.9853 | -103.0897
20.989 | -103.1027
20.9927 | -103.1151
20.9965 | -103.1284
21.0002 | -103.1408
21.004 | -103.1534
21.0077 | -103.1679
21.0115 | -103.1844
21.0152 | -103.1982
21.0189 | -103.2112
21.0227 | -103.2246
21.0264 | -103.2363
21.0302 | -103.246
21.0339 | -103.2567
21.0377 | -103.2634
21.0414 | -103.269
21.0451 | -103.2751
21.0489 | -103.2822
21.0526 | -103.2896
21.0564 | -103.2923
21.0601 | -103.2974
21.0639 | -103.3048
21.0676 | -103.3116
21.0713 | -103.3161
21.0751 | -103.3188
21.0788 | -103.3214
21.0826 | -103.325
21.0863 | -103.3271
21.0901 | -103.3251
21.0938 | -103.3242
21.0975 | -103.3238
21.1013 | -103.3248
21.105 | -103.3231
21.1088 | -103.3203
21.1125 | -103.3168
21.1163 | -103.3128
21.12 | -103.3096
21.1237 | -103.3071
21.1275 | -103.3033
21.1312 | -103.2983
21.135 | -103.2951
21.1387 | -103.292
21.1425 | -103.2864
21.1462 | -103.2819
21.1499 | -103.2763
21.1537 | -103.272
21.1574 | -103.2672
21.1612 | -103.2614
21.1649 | -103.2559
21.1687 | -103.2516
21.1724 | -103.2468
21.1761 | -103.2403
21.1799 | -103.2332
21.1836 | -103.2259
21.1874 | -103.2196
21.1911 | -103.2145
21.1949 | -103.2114
21.1986 | -103.2084
21.2023 | -103.205
21.2061 | -103.2021
21.2098 | -103.2014
21.2136 | -103.2008
21.2173 | -103.1998
21.2211 | -103.1992
21.2248 | -103.1985
21.2285 | -103.1986
21.2323 | -103.1979
21.236 | -103.1971
21.2398 | -103.1999
21.2435 | -103.2021
21.2473 | -103.2051
21.251 | -103.2083
21.2547 | -103.2098
21.2585 | -103.2117
21.2622 | -103.2145
21.266 | -103.2178
21.2697 | -103.222
21.2735 | -103.2247
21.2772 | -103.2289
21.2809 | -103.2338
21.2847 | -103.2381
21.2884 | -103.2426
21.2922 | -103.245
21.2959 | -103.2465
21.2996 | -103.2502
21.3034 | -103.253
21.3071 | -103.2533
21.3109 | -103.2534
21.3146 | -103.2527
21.3184 | -103.2519
21.3221 | -103.251
21.3258 | -103.2507
21.3296 | -103.2527
21.3333 | -103.2536
21.3371 | -103.2539
21.3408 | -103.2545
21.3446 | -103.2523
21.3483 | -103.2473
21.352 | -103.2431
21.3558 | -103.2375
21.3595 | -103.2306
21.3633 | -103.2251
21.367 | -103.2193
21.3708 | -103.2138
21.3745 | -103.2087
21.3782 | -103.2024
21.382 | -103.1958
21.3857 | -103.1912
21.3895 | -103.1867
21.3932 | -103.1813
21.397 | -103.1745
21.4007 | -103.1675
21.4044 | -103.1597
21.4082 | -103.1518
21.4119 | -103.1443
21.4157 | -103.1353
21.4194 | -103.1255
21.4232 | -103.1151
21.4269 | -103.105
21.4306 | -103.0965
21.4344 | -103.0894
21.4381 | -103.0822
21.4419 | -103.0755
21.4456 | -103.0717
21.4494 | -103.065
21.4531 | -103.0565
21.4568 | -103.047
21.4606 | -103.0382
21.4643 | -103.0303
21.4681 | -103.0208
21.4718 | -103.0118
21.4756 | -103.0031
21.4793 | -102.9964
21.483 | -102.9912
21.4868 | -102.9861
21.4905 | -102.9789
21.4943 | -102.97
21.498 | -102.9576
21.5018 | -102.9469
21.5055 | -102.9376
21.5092 | -102.9269
21.513 | -102.9147
21.5167 | -102.9022
21.5205 | -102.8872
21.5242 | -102.8702
21.528 | -102.8539
21.5317 | -102.8385
21.5354 | -102.8239
21.5392 | -102.8098
21.5429 | -102.7942
21.5467 | -102.7757
21.5504 | -102.7564
21.5542 | -102.7383
21.5579 | -102.721
21.5616 | -102.7028
21.5654 | -102.6874
21.5691 | -102.6724
21.5729 | -102.6572
21.5766 | -102.6427
21.5804 | -102.6291
21.5841 | -102.614
21.5878 | -102.6
21.5916 | -102.5878
21.5953 | -102.5851
21.5991 | -102.578
21.6028 | -102.5721
21.6065 | -102.5661
21.6103 | -102.5615
21.614 | -102.56
21.6178 | -102.5601
21.6215 | -102.5612
21.6253 | -102.5621
21.629 | -102.5643
21.6327 | -102.5662
21.6365 | -102.5692
21.6402 | -102.5748
21.644 | -102.5808
21.6477 | -102.588
21.6515 | -102.597
21.6552 | -102.6081
21.6589 | -102.619
21.6627 | -102.6311
21.6664 | -102.6472
21.6702 | -102.6624
21.6739 | -102.6755
21.6777 | -102.6881
21.6814 | -102.7015
21.6851 | -102.7169
21.6889 | -102.7337
21.6926 | -102.7494
21.6964 | -102.7635
21.7001 | -102.7789
21.7039 | -102.7948
21.7076 | -102.8098
21.7113 | -102.8251
21.7151 | -102.8403
21.7188 | -102.8564
21.7226 | -102.8711
21.7263 | -102.8832
21.7301 | -102.8961
21.7338 | -102.9082
21.7375 | -102.9165
21.7413 | -102.919
21.745 | -102.9205
21.7488 | -102.9214
21.7525 | -102.9177
21.7563 | -102.917
21.76 | -102.9178
21.7637 | -102.9166
21.7675 | -102.9116
21.7712 | -102.902
21.775 | -102.8927
21.7787 | -102.8832
21.7825 | -102.8722
21.7862 | -102.8572
21.7899 | -102.8375
21.7937 | -102.8149
21.7974 | -102.7911
21.8012 | -102.7691
21.8049 | -102.7479
21.8087 | -102.7251
21.8124 | -102.7025
21.8161 | -102.6782
21.8199 | -102.6515
21.8236 | -102.6256
21.8274 | -102.6016
21.8311 | -102.5785
21.8349 | -102.5536
21.8386 | -102.5267
21.8423 | -102.5016
21.8461 | -102.4749
21.8498 | -102.4454
21.8536 | -102.4187
21.8573 | -102.3957
21.8611 | -102.3713
21.8648 | -102.3477
21.8685 | -102.3263
21.8723 | -102.3023
21.876 | -102.2781
21.8798 | -102.2557
21.8835 | -102.2353
21.8873 | -102.2164
21.891 | -102.2006
21.8947 | -102.186
21.8985 | -102.1698
21.9022 | -102.1569
21.906 | -102.1449
21.9097 | -102.1356
21.9135 | -102.1267
21.9172 | -102.1178
21.9209 | -102.1103
21.9247 | -102.1044
21.9284 | -102.0987
21.9322 | -102.0952
21.9359 | -102.0929
21.9396 | -102.0917
21.9434 | -102.0918
21.9471 | -102.0925
21.9509 | -102.0934
21.9546 | -102.0956
21.9584 | -102.0991
21.9621 | -102.1009
21.9658 | -102.106
21.9696 | -102.1133
21.9733 | -102.1206
21.9771 | -102.1272
21.9808 | -102.1335
21.9846 | -102.1406
21.9883 | -102.1477
21.992 | -102.1554
21.9958 | -102.1609
21.9995 | -102.1644
22.0033 | -102.1685
22.007 | -102.1754
22.0108 | -102.1784
22.0145 | -102.1811
22.0182 | -102.1846
22.022 | -102.1875
22.0257 | -102.1906
22.0295 | -102.1933
22.0332 | -102.1957
22.037 | -102.1973
22.0407 | -102.1991
22.0444 | -102.1995
22.0482 | -102.1994
22.0519 | -102.2014
22.0557 | -102.2023
22.0594 | -102.2026
22.0632 | -102.2033
22.0669 | -102.2012
22.0706 | -102.1979
22.0744 | -102.1942
22.0781 | -102.1897
22.0819 | -102.1842
22.0856 | -102.1762
22.0894 | -102.1649
22.0931 | -102.1529
22.0968 | -102.1409
22.1006 | -102.1309
22.1043 | -102.1217
22.1081 | -102.1131
22.1118 | -102.1032
22.1156 | -102.0919
22.1193 | -102.0801
22.123 | -102.0667
22.1268 | -102.0539
22.1305 | -102.0424
22.1343 | -102.0301
22.138 | -102.0198
22.1418 | -102.0083
22.1455 | -101.9953
22.1492 | -101.9839
22.153 | -101.9745
22.1567 | -101.9645
22.1605 | -101.9557
22.1642 | -101.9479
22.168 | -101.9419
22.1717 | -101.9352
22.1754 | -101.9296
22.1792 | -101.9242
22.1829 | -101.918
22.1867 | -101.9168
22.1904 | -101.9167
22.1942 | -101.9165
22.1979 | -101.9169
22.2016 | -101.9176
22.2054 | -101.9201
22.2091 | -101.9232
22.2129 | -101.9284
22.2166 | -101.9326
22.2204 | -101.9374
22.2241 | -101.9412
22.2278 | -101.9475
22.2316 | -101.9568
22.2353 | -101.9668
22.2391 | -101.9765
22.2428 | -101.9871
22.2465 | -101.9978
22.2503 | -102.0096
22.254 | -102.0234
22.2578 | -102.0392
22.2615 | -102.0562
22.2653 | -102.0725
22.269 | -102.088
22.2727 | -102.1025
22.2765 | -102.1179
22.2802 | -102.1309
22.284 | -102.1429
22.2877 | -102.1549
22.2915 | -102.1671
22.2952 | -102.1817
22.2989 | -102.1961
22.3027 | -102.2108
22.3064 | -102.2255
22.3102 | -102.2412
22.3139 | -102.259
22.3177 | -102.2743
22.3214 | -102.2883
22.3251 | -102.3034
22.3289 | -102.3158
22.3326 | -102.3274
22.3364 | -102.3383
22.3401 | -102.3486
22.3439 | -102.3575
22.3476 | -102.3676
22.3513 | -102.3764
22.3551 | -102.3831
22.3588 | -102.3888
22.3626 | -102.3961
22.3663 | -102.4058
22.3701 | -102.416
22.3738 | -102.4243
22.3775 | -102.4328
22.3813 | -102.4405
22.385 | -102.4479
22.3888 | -102.4561
22.3925 | -102.4646
22.3963 | -102.475
22.4 | -102.4861
22.4037 | -102.4962
22.4075 | -102.5064
22.4112 | -102.5164
22.415 | -102.5252
22.4187 | -102.5321
22.4225 | -102.5373
22.4262 | -102.5403
22.4299 | -102.5449
22.4337 | -102.5488
22.4374 | -102.5517
22.4412 | -102.5532
22.4449 | -102.5548
22.4487 | -102.5575
22.4524 | -102.5593
22.4561 | -102.5595
22.4599 | -102.5599
22.4636 | -102.5595
22.4674 | -102.5528
22.4711 | -102.5441
22.4749 | -102.5349
22.4786 | -102.5259
22.4823 | -102.5198
22.4861 | -102.5102
22.4898 | -102.501
22.4936 | -102.4921
22.4973 | -102.481
22.5011 | -102.4689
22.5048 | -102.4535
22.5085 | -102.4372
22.5123 | -102.4214
22.516 | -102.4066
22.5198 | -102.3904
22.5235 | -102.3721
22.5273 | -102.3518
22.531 | -102.3314
22.5347 | -102.3095
22.5385 | -102.2868
22.5422 | -102.2648
22.546 | -102.2395
22.5497 | -102.2113
22.5535 | -102.1817
22.5572 | -102.1505
22.5609 | -102.1191
22.5647 | -102.0869
22.5684 | -102.0536
22.5722 | -102.0188
22.5759 | -101.9838
22.5796 | -101.9481
22.5834 | -101.9116
22.5871 | -101.876
22.5909 | -101.8417
22.5946 | -101.809
22.5984 | -101.7769
22.6021 | -101.7434
22.6058 | -101.7125
22.6096 | -101.6836
22.6133 | -101.6549
22.6171 | -101.6281
22.6208 | -101.6014
22.6246 | -101.5754
22.6283 | -101.551
22.632 | -101.5252
22.6358 | -101.501
22.6395 | -101.4786
22.6433 | -101.4616
22.647 | -101.4472
22.6508 | -101.4338
22.6545 | -101.4218
22.6582 | -101.4107
22.662 | -101.4018
22.6657 | -101.3929
22.6695 | -101.387
22.6732 | -101.3828
22.677 | -101.3874
22.6807 | -101.3881
22.6844 | -101.3891
22.6882 | -101.3899
22.6919 | -101.3898
22.6957 | -101.391
22.6994 | -101.3929
22.7032 | -101.3955
22.7069 | -101.4001
22.7106 | -101.4065
22.7144 | -101.4142
22.7181 | -101.4212
22.7219 | -101.4281
22.7256 | -101.4349
22.7294 | -101.4413
22.7331 | -101.4488
22.7368 | -101.4581
22.7406 | -101.4696
22.7443 | -101.4815
22.7481 | -101.4943
22.7518 | -101.5075
22.7556 | -101.5202
22.7593 | -101.534
22.763 | -101.5492
22.7668 | -101.5632
22.7705 | -101.5762
22.7743 | -101.5904
22.778 | -101.6064
22.7818 | -101.6226
22.7855 | -101.6396
22.7892 | -101.6586
22.793 | -101.6742
22.7967 | -101.689
22.8005 | -101.7028
22.8042 | -101.7135
22.808 | -101.726
22.8117 | -101.7378
22.8154 | -101.7496
22.8192 | -101.7635
22.8229 | -101.7785
22.8267 | -101.794
22.8304 | -101.8096
22.8342 | -101.8241
22.8379 | -101.8363
22.8416 | -101.8467
22.8454 | -101.8559
22.8491 | -101.8665
22.8529 | -101.8756
22.8566 | -101.8833
22.8604 | -101.8899
22.8641 | -101.8958
22.8678 | -101.9021
22.8716 | -101.9089
22.8753 | -101.9145
22.8791 | -101.9196
22.8828 | -101.9228
22.8865 | -101.9246
22.8903 | -101.9236
22.894 | -101.9229
22.8978 | -101.9185
22.9015 | -101.9142
22.9053 | -101.912
22.909 | -101.9096
22.9127 | -101.906
22.9165 | -101.9023
22.9202 | -101.8981
22.924 | -101.8948
22.9277 | -101.8896
22.9315 | -101.8826
22.9352 | -101.8774
22.9389 | -101.8727
22.9427 | -101.8673
22.9464 | -101.8606
22.9502 | -101.8544
22.9539 | -101.8513
22.9577 | -101.848
22.9614 | -101.8456
22.9651 | -101.8425
22.9689 | -101.8389
22.9726 | -101.8352
22.9764 | -101.8346
22.9801 | -101.8363
22.9839 | -101.838
22.9876 | -101.8407
22.9913 | -101.8444
22.9951 | -101.8487
22.9988 | -101.8536
23.0026 | -101.8575
23.0063 | -101.8601
23.0101 | -101.8639
23.0138 | -101.8687
23.0175 | -101.8761
23.0213 | -101.8816
23.025 | -101.8871
23.0288 | -101.8918
23.0325 | -101.895
23.0363 | -101.8979
23.04 | -101.9025
23.0437 | -101.9065
23.0475 | -101.9118
23.0512 | -101.918
23.055 | -101.923
23.0587 | -101.925
23.0625 | -101.9258
23.0662 | -101.9245
23.0699 | -101.9221
23.0737 | -101.9196
23.0774 | -101.914
23.0812 | -101.9077
23.0849 | -101.9028
23.0887 | -101.8979
23.0924 | -101.8924
23.0961 | -101.8865
23.0999 | -101.8795
23.1036 | -101.871
23.1074 | -101.86
23.1111 | -101.8487
23.1149 | -101.8385
23.1186 | -101.8289
23.1223 | -101.8183
23.1261 | -101.8067
23.1298 | -101.7948
23.1336 | -101.7813
23.1373 | -101.7664
23.1411 | -101.7516
23.1448 | -101.7364
23.1485 | -101.7172
23.1523 | -101.697
23.156 | -101.6778
23.1598 | -101.6595
23.1635 | -101.6416
23.1673 | -101.6267
23.171 | -101.6105
23.1747 | -101.5915
23.1785 | -101.5738
23.1822 | -101.5576
23.186 | -101.5402
23.1897 | -101.5225
23.1934 | -101.5051
23.1972 | -101.4867
23.2009 | -101.4695
23.2047 | -101.4531
23.2084 | -101.437
23.2122 | -101.4219
23.2159 | -101.4099
23.2196 | -101.3966
23.2234 | -101.384
23.2271 | -101.3746
23.2309 | -101.3645
23.2346 | -101.3539
23.2384 | -101.344
23.2421 | -101.3353
23.2458 | -101.3292
23.2496 | -101.3238
23.2533 | -101.3191
23.2571 | -101.3154
23.2608 | -101.313
23.2646 | -101.3115
23.2683 | -101.3116
23.272 | -101.3147
23.2758 | -101.3189
23.2795 | -101.3236
23.2833 | -101.3289
23.287 | -101.3374
23.2908 | -101.3459
23.2945 | -101.3547
23.2982 | -101.3642
23.302 | -101.373
23.3057 | -101.3814
23.3095 | -101.3906
23.3132 | -101.4008
23.317 | -101.4128
23.3207 | -101.4269
23.3244 | -101.4415
23.3282 | -101.4571
23.3319 | -101.4727
23.3357 | -101.4879
23.3394 | -101.5034
23.3432 | -101.5192
23.3469 | -101.5337
23.3506 | -101.5499
23.3544 | -101.5678
23.3581 | -101.5868
23.3619 | -101.604
23.3656 | -101.6184
23.3694 | -101.6306
23.3731 | -101.6434
23.3768 | -101.6545
23.3806 | -101.6667
23.3843 | -101.6804
23.3881 | -101.6956
23.3918 | -101.7102
23.3956 | -101.724
23.3993 | -101.7386
23.403 | -101.7529
23.4068 | -101.7686
23.4105 | -101.7844
23.4143 | -101.7989
23.418 | -101.8127
23.4218 | -101.8259
23.4255 | -101.8385
23.4292 | -101.8515
23.433 | -101.8645
23.4367 | -101.8786
23.4405 | -101.8931
23.4442 | -101.9079
23.448 | -101.924
23.4517 | -101.939
23.4554 | -101.9539
23.4592 | -101.9685
23.4629 | -101.9819
23.4667 | -101.9953
23.4704 | -102.0085
23.4742 | -102.0209
23.4779 | -102.0373
23.4816 | -102.0533
23.4854 | -102.0662
23.4891 | -102.0775
23.4929 | -102.091
23.4966 | -102.1037
23.5004 | -102.1184
23.5041 | -102.1324
23.5078 | -102.1464
23.5116 | -102.1597
23.5153 | -102.1725
23.5191 | -102.1815
23.5228 | -102.1879
23.5266 | -102.1951
23.5303 | -102.2014
23.534 | -102.2068
23.5378 | -102.2122
23.5415 | -102.2164
23.5453 | -102.219
23.549 | -102.2225
23.5527 | -102.2238
23.5565 | -102.222
23.5602 | -102.2176
23.564 | -102.2112
23.5677 | -102.206
23.5715 | -102.1997
23.5752 | -102.1938
23.5789 | -102.1825
23.5827 | -102.172
23.5864 | -102.1618
23.5902 | -102.1534
23.5939 | -102.1429
23.5977 | -102.1317
23.6014 | -102.1199
23.6051 | -102.1075
23.6089 | -102.0951
23.6126 | -102.0794
23.6164 | -102.0639
23.6201 | -102.0483
23.6239 | -102.0322
23.6276 | -102.0155
23.6313 | -102.0015
23.6351 | -101.9896
23.6388 | -101.9752
23.6426 | -101.9601
23.6463 | -101.9455
23.6501 | -101.9299
23.6538 | -101.9164
23.6575 | -101.9041
23.6613 | -101.8911
23.665 | -101.8805
23.6688 | -101.8723
23.6725 | -101.8641
23.6763 | -101.8543
23.68 | -101.8456
23.6837 | -101.838
23.6875 | -101.8306
23.6912 | -101.8243
23.695 | -101.82
23.6987 | -101.8167
23.7025 | -101.8144
23.7062 | -101.8123
23.7099 | -101.8111
23.7137 | -101.8077
23.7174 | -101.8024
23.7212 | -101.798
23.7249 | -101.7944
23.7287 | -101.792
23.7324 | -101.7903
23.7361 | -101.7895
23.7399 | -101.789
23.7436 | -101.7883
23.7474 | -101.7858
23.7511 | -101.7821
23.7549 | -101.7708
23.7586 | -101.7662
23.7623 | -101.7621
23.7661 | -101.7573
23.7698 | -101.752
23.7736 | -101.7472
23.7773 | -101.7423
23.7811 | -101.7352
23.7848 | -101.7259
23.7885 | -101.7171
23.7923 | -101.7081
23.796 | -101.7006
23.7998 | -101.6942
23.8035 | -101.6865
23.8073 | -101.6772
23.811 | -101.6689
23.8147 | -101.6614
23.8185 | -101.6541
23.8222 | -101.6498
23.826 | -101.643
23.8297 | -101.6304
23.8334 | -101.6179
23.8372 | -101.6074
23.8409 | -101.5998
23.8447 | -101.5907
23.8484 | -101.5791
23.8522 | -101.5676
23.8559 | -101.5606
23.8596 | -101.5545
23.8634 | -101.5459
23.8671 | -101.539
23.8709 | -101.5335
23.8746 | -101.5287
23.8784 | -101.5237
23.8821 | -101.5203
23.8858 | -101.5167
23.8896 | -101.5141
23.8933 | -101.5125
23.8971 | -101.5114
23.9008 | -101.5105
23.9046 | -101.5075
23.9083 | -101.5051
23.912 | -101.5029
23.9158 | -101.5024
23.9195 | -101.501
23.9233 | -101.5007
23.927 | -101.4994
23.9308 | -101.4977
23.9345 | -101.4961
23.9382 | -101.4954
23.942 | -101.4955
23.9457 | -101.4957
23.9495 | -101.4988
23.9532 | -101.5021
23.957 | -101.5041
23.9607 | -101.5057
23.9644 | -101.5076
23.9682 | -101.51
23.9719 | -101.5105
23.9757 | -101.5115
23.9794 | -101.5137
23.9832 | -101.5158
23.9869 | -101.5185
23.9906 | -101.5182
23.9944 | -101.5196
23.9981 | -101.5215
24.0019 | -101.5243
24.0056 | -101.526
24.0094 | -101.528
24.0131 | -101.5322
24.0168 | -101.538
24.0206 | -101.5435
24.0243 | -101.5484
24.0281 | -101.5528
24.0318 | -101.5578
24.0356 | -101.5634
24.0393 | -101.5702
24.043 | -101.5761
24.0468 | -101.5816
24.0505 | -101.588
24.0543 | -101.5949
24.058 | -101.597
24.0618 | -101.5973
24.0655 | -101.5988
24.0692 | -101.5995
24.073 | -101.5999
24.0767 | -101.6005
24.0805 | -101.6001
24.0842 | -101.6013
24.088 | -101.6021
24.0917 | -101.6033
24.0954 | -101.6035
24.0992 | -101.5998
24.1029 | -101.5948
24.1067 | -101.5901
24.1104 | -101.5847
24.1142 | -101.5791
24.1179 | -101.5726
24.1216 | -101.5644
24.1254 | -101.5558
24.1291 | -101.5472
24.1329 | -101.5379
24.1366 | -101.5248
24.1404 | -101.5112
24.1441 | -101.5001
24.1478 | -101.491
24.1516 | -101.4804
24.1553 | -101.4691
24.1591 | -101.4582
24.1628 | -101.4458
24.1666 | -101.4363
24.1703 | -101.4291
24.174 | -101.4215
24.1778 | -101.4107
24.1815 | -101.3976
24.1853 | -101.3866
24.189 | -101.3799
24.1927 | -101.3748
24.1965 | -101.3686
24.2002 | -101.3621
24.204 | -101.3573
24.2077 | -101.3514
24.2115 | -101.3437
24.2152 | -101.3361
24.2189 | -101.3302
24.2227 | -101.3249
24.2264 | -101.3188
24.2302 | -101.3113
24.2339 | -101.3044
24.2377 | -101.2996
24.2414 | -101.2974
24.2451 | -101.2991
24.2489 | -101.3026
24.2526 | -101.3028
24.2564 | -101.3029
24.2601 | -101.298
24.2639 | -101.2742
24.2676 | -101.2186
24.2713 | -101.1299
24.2751 | -101.0473
24.2788 | -100.9606
24.2826 | -100.8753
24.2863 | -100.7952
24.2901 | -100.7237
24.2938 | -100.6589
24.2975 | -100.539
24.3013 | -100.4645
24.305 | -100.3872
24.3088 | -100.305
24.3125 | -100.2189
24.3163 | -100.1308
24.32 | -100.0381
24.3237 | -99.9409
24.3275 | -99.8415
24.3312 | -99.7391
24.335 | -99.634
24.3387 | -99.5254
24.3425 | -99.4151
24.3462 | -99.3039
24.3499 | -99.193
24.3537 | -99.0815
24.3574 | -98.9698
24.3612 | -98.858
24.3649 | -98.7472
24.3687 | -98.6381
24.3724 | -98.5316
24.3761 | -98.4277
24.3799 | -98.3274
24.3836 | -98.2297
24.3874 | -98.1361
24.3911 | -98.0469
24.3949 | -97.9626
24.3986 | -97.8834
24.4023 | -97.8102
24.4061 | -97.7422
24.4098 | -97.6789
24.4136 | -97.6211
24.4173 | -97.569
24.4211 | -97.5238
24.4248 | -97.4857
24.4285 | -97.4541
24.4323 | -97.4298
24.436 | -97.4115
24.4398 | -97.3995
24.4435 | -97.395
24.4473 | -97.3973
24.451 | -97.4061
24.4547 | -97.4217
24.4585 | -97.4439
24.4622 | -97.4727
24.466 | -97.5088
24.4697 | -97.5514
24.4734 | -97.5997
24.4772 | -97.6541
24.4809 | -97.7146
24.4847 | -97.7805
24.4884 | -97.852
24.4922 | -97.9277
24.4959 | -98.0084
24.4996 | -98.0939
24.5034 | -98.1833
24.5071 | -98.2774
24.5109 | -98.3742
24.5146 | -98.4746
24.5184 | -98.5781
24.5221 | -98.6852
24.5258 | -98.7958
24.5296 | -98.9075
24.5333 | -99.0186
24.5371 | -99.1298
24.5408 | -99.2406
24.5446 | -99.353
24.5483 | -99.464
24.552 | -99.5732
24.5558 | -99.6818
24.5595 | -99.7891
24.5633 | -99.8894
24.567 | -99.9813
24.5708 | -100.0581
24.5745 | -100.1263
24.5782 | -100.1933
24.582 | -100.2575
24.5857 | -100.3167
24.5895 | -100.3736
24.5932 | -100.425
24.597 | -100.4725
24.6007 | -100.514
24.6044 | -100.5529
24.6082 | -100.5982
24.6119 | -100.6569
24.6157 | -100.7265
24.6194 | -100.7705
24.6232 | -100.8009
24.6269 | -100.814
24.6306 | -100.8094
24.6344 | -100.7841
24.6381 | -100.7434
24.6419 | -100.701
24.6456 | -100.6572
24.6494 | -100.6098
24.6531 | -100.5597
24.6568 | -100.5103
24.6606 | -100.4585
24.6643 | -100.4052
24.6681 | -100.3531
24.6718 | -100.3006
24.6756 | -100.2454
24.6793 | -100.1915
24.683 | -100.1405
24.6868 | -100.0897
24.6905 | -100.0414
24.6943 | -99.9939
24.698 | -99.9485
24.7018 | -99.9075
24.7055 | -99.8691
24.7092 | -99.8337
24.713 | -99.8013
24.7167 | -99.7714
24.7205 | -99.7427
24.7242 | -99.7159
24.728 | -99.6924
24.7317 | -99.673
24.7354 | -99.6581
24.7392 | -99.647
24.7429 | -99.6389
24.7467 | -99.6351
24.7504 | -99.6346
24.7542 | -99.6373
24.7579 | -99.6436
24.7616 | -99.6535
24.7654 | -99.6664
24.7691 | -99.6807
24.7729 | -99.6983
24.7766 | -99.7181
24.7804 | -99.7418
24.7841 | -99.7688
24.7878 | -99.7986
24.7916 | -99.8304
24.7953 | -99.8629
24.7991 | -99.8988
24.8028 | -99.9365
24.8066 | -99.9757
24.8103 | -100.0144
24.814 | -100.0541
24.8178 | -100.0954
24.8215 | -100.1356
24.8253 | -100.1751
24.829 | -100.2136
24.8327 | -100.251
24.8365 | -100.2757
24.8402 | -100.3076
24.844 | -100.3417
24.8477 | -100.3755
24.8515 | -100.4067
24.8552 | -100.4379
24.8589 | -100.4669
24.8627 | -100.4951
24.8664 | -100.5213
24.8702 | -100.5456
24.8739 | -100.5649
24.8777 | -100.5814
24.8814 | -100.5953
24.8851 | -100.6084
24.8889 | -100.6157
24.8926 | -100.6211
24.8964 | -100.6236
24.9001 | -100.625
24.9039 | -100.6257
24.9076 | -100.6241
24.9113 | -100.6206
24.9151 | -100.6189
24.9188 | -100.6198
24.9226 | -100.6298
24.9263 | -100.6378
24.9301 | -100.6356
24.9338 | -100.6263
24.9375 | -100.613
24.9413 | -100.5905
24.945 | -100.5666
24.9488 | -100.5427
24.9525 | -100.5198
24.9563 | -100.4979
24.96 | -100.475
24.9637 | -100.4502
24.9675 | -100.4268
24.9712 | -100.4044
24.975 | -100.3816
24.9787 | -100.3594
24.9825 | -100.3367
24.9862 | -100.3135
24.9899 | -100.2923
24.9937 | -100.2725
24.9974 | -100.2546
25.0012 | -100.2375
25.0049 | -100.2216
25.0087 | -100.2061
25.0124 | -100.1914
25.0161 | -100.1786
25.0199 | -100.1676
25.0236 | -100.1592
25.0274 | -100.1532
25.0311 | -100.1466
25.0349 | -100.1376
25.0386 | -100.1316
25.0423 | -100.1269
25.0461 | -100.1225
25.0498 | -100.1185
25.0536 | -100.1139
25.0573 | -100.1117
25.0611 | -100.1107
25.0648 | -100.1103
25.0685 | -100.1105
25.0723 | -100.1139
25.076 | -100.1208
25.0798 | -100.1283
25.0835 | -100.1363
25.0873 | -100.1443
25.091 | -100.1533
25.0947 | -100.1645
25.0985 | -100.1753
25.1022 | -100.1881
25.106 | -100.2022
25.1097 | -100.2171
25.1135 | -100.2337
25.1172 | -100.2521
25.1209 | -100.2716
25.1247 | -100.2922
25.1284 | -100.314
25.1322 | -100.3336
25.1359 | -100.3526
25.1396 | -100.3733
25.1434 | -100.3949
25.1471 | -100.4186
25.1509 | -100.4399
25.1546 | -100.462
25.1584 | -100.4841
25.1621 | -100.5089
25.1658 | -100.5347
25.1696 | -100.5598
25.1733 | -100.5848
25.1771 | -100.6121
25.1808 | -100.6404
25.1846 | -100.6669
25.1883 | -100.6912
25.192 | -100.7162
25.1958 | -100.7423
25.1995 | -100.7641
25.2033 | -100.7851
25.207 | -100.8072
25.2108 | -100.8266
25.2145 | -100.8441
25.2182 | -100.8625
25.222 | -100.8811
25.2257 | -100.899
25.2295 | -100.9191
25.2332 | -100.9378
25.237 | -100.9536
25.2407 | -100.97
25.2444 | -100.9901
25.2482 | -101.009
25.2519 | -101.0259
25.2557 | -101.0403
25.2594 | -101.0531
25.2632 | -101.0656
25.2669 | -101.0777
25.2706 | -101.0896
25.2744 | -101.0986
25.2781 | -101.1058
25.2819 | -101.1131
25.2856 | -101.1189
25.2894 | -101.1227
25.2931 | -101.1249
25.2968 | -101.1285
25.3006 | -101.1311
25.3043 | -101.133
25.3081 | -101.1307
25.3118 | -101.1274
25.3156 | -101.1284
25.3193 | -101.1322
25.323 | -101.1345
25.3268 | -101.1358
25.3305 | -101.1352
25.3343 | -101.1332
25.338 | -101.1307
25.3418 | -101.1299
25.3455 | -101.1289
25.3492 | -101.126
25.353 | -101.1213
25.3567 | -101.1164
25.3605 | -101.1106
25.3642 | -101.1045
25.368 | -101.1007
25.3717 | -101.0976
25.3754 | -101.0953
25.3792 | -101.0892
25.3829 | -101.0833
25.3867 | -101.0782
25.3904 | -101.073
25.3942 | -101.0695
25.3979 | -101.0661
25.4016 | -101.0609
25.4054 | -101.0563
25.4091 | -101.0535
25.4129 | -101.0503
25.4166 | -101.0445
25.4204 | -101.0403
25.4241 | -101.0374
25.4278 | -101.0352
25.4316 | -101.032
25.4353 | -101.0277
25.4391 | -101.0218
25.4428 | -101.0156
25.4465 | -101.0098
25.4503 | -101.0022
25.454 | -100.9939
25.4578 | -100.9894
25.4615 | -100.9856
25.4653 | -100.9804
25.469 | -100.9767
25.4727 | -100.9712
25.4765 | -100.9637
25.4802 | -100.9573
25.484 | -100.9496
25.4877 | -100.9414
25.4915 | -100.9361
25.4952 | -100.9323
25.4989 | -100.9263
25.5027 | -100.9206
25.5064 | -100.9149
25.5102 | -100.9093
25.5139 | -100.9019
25.5177 | -100.8951
25.5214 | -100.8886
25.5251 | -100.8821
25.5289 | -100.8764
25.5326 | -100.872
25.5364 | -100.8649
25.5401 | -100.8585
25.5439 | -100.8538
25.5476 | -100.849
25.5513 | -100.8432
25.5551 | -100.8394
25.5588 | -100.836
25.5626 | -100.8307
25.5663 | -100.8252
25.5701 | -100.8204
25.5738 | -100.8171
25.5775 | -100.8138
25.5813 | -100.8091
25.585 | -100.803
25.5888 | -100.7967
25.5925 | -100.7931
25.5963 | -100.7908
25.6 | -100.7876
25.6037 | -100.7846
25.6075 | -100.7806
25.6112 | -100.7782
25.615 | -100.7763
25.6187 | -100.7743
25.6225 | -100.7723
25.6262 | -100.7675
25.6299 | -100.7628
25.6337 | -100.7602
25.6374 | -100.7575
25.6412 | -100.7563
25.6449 | -100.7556
25.6487 | -100.7547
25.6524 | -100.7541
25.6561 | -100.7516
25.6599 | -100.7451
25.6636 | -100.7387
25.6674 | -100.7349
25.6711 | -100.7278
25.6749 | -100.7169
25.6786 | -100.7072
25.6823 | -100.6986
25.6861 | -100.688
25.6898 | -100.6776
25.6936 | -100.6687
25.6973 | -100.659
25.7011 | -100.6486
25.7048 | -100.636
25.7085 | -100.6246
25.7123 | -100.614
25.716 | -100.6038
25.7198 | -100.5938
25.7235 | -100.582
25.7273 | -100.5656
25.731 | -100.5489
25.7347 | -100.5295
25.7385 | -100.5104
25.7422 | -100.4936
25.746 | -100.475
25.7497 | -100.4568
25.7535 | -100.4418
25.7572 | -100.4235
25.7609 | -100.404
25.7647 | -100.3867
25.7684 | -100.3677
25.7722 | -100.3468
25.7759 | -100.3271
25.7796 | -100.3068
25.7834 | -100.287
25.7871 | -100.2676
25.7909 | -100.249
25.7946 | -100.2302
25.7984 | -100.2113
25.8021 | -100.1931
25.8058 | -100.1752
25.8096 | -100.156
25.8133 | -100.1392
25.8171 | -100.1242
25.8208 | -100.1086
25.8246 | -100.0893
25.8283 | -100.0689
25.832 | -100.0493
25.8358 | -100.0322
25.8395 | -100.0156
25.8433 | -99.9984
25.847 | -99.9821
25.8508 | -99.9679
25.8545 | -99.9535
25.8582 | -99.939
25.862 | -99.9255
25.8657 | -99.9122
25.8695 | -99.8986
25.8732 | -99.8869
25.877 | -99.8763
25.8807 | -99.8653
25.8844 | -99.8547
25.8882 | -99.8426
25.8919 | -99.8303
25.8957 | -99.8189
25.8994 | -99.8081
25.9032 | -99.7973
25.9069 | -99.7877
25.9106 | -99.7793
25.9144 | -99.7724
25.9181 | -99.7645
25.9219 | -99.7555
25.9256 | -99.7466
25.9294 | -99.739
25.9331 | -99.7311
25.9368 | -99.7237
25.9406 | -99.7154
25.9443 | -99.7065
25.9481 | -99.6993
25.9518 | -99.6918
25.9556 | -99.6864
25.9593 | -99.6807
25.963 | -99.6762
25.9668 | -99.6729
25.9705 | -99.6701
25.9743 | -99.6678
25.978 | -99.6654
25.9818 | -99.6647
25.9855 | -99.663
25.9892 | -99.6618
25.993 | -99.6617
25.9967 | -99.6632
26.0005 | -99.6662
26.0042 | -99.6695
26.008 | -99.6714
26.0117 | -99.6733
26.0154 | -99.6746
26.0192 | -99.6746
26.0229 | -99.6791
26.0267 | -99.6872
26.0304 | -99.6958
26.0342 | -99.7038
26.0379 | -99.7124
26.0416 | -99.72
26.0454 | -99.7254
26.0491 | -99.7297
26.0529 | -99.7354
26.0566 | -99.7398
26.0604 | -99.7418
26.0641 | -99.7445
26.0678 | -99.7471
26.0716 | -99.7493
26.0753 | -99.7493
26.0791 | -99.747
26.0828 | -99.744
26.0865 | -99.7417
26.0903 | -99.7407
26.094 | -99.7404
26.0978 | -99.7382
26.1015 | -99.7372
26.1053 | -99.7353
26.109 | -99.7305
26.1127 | -99.7247
26.1165 | -99.7172
26.1202 | -99.7063
26.124 | -99.6968
26.1277 | -99.6835
26.1315 | -99.6691
26.1352 | -99.6568
26.1389 | -99.6446
26.1427 | -99.6305
26.1464 | -99.6178
26.1502 | -99.6039
26.1539 | -99.5887
26.1577 | -99.5749
26.1614 | -99.5622
26.1651 | -99.5496
26.1689 | -99.5367
26.1726 | -99.5247
26.1764 | -99.5139
26.1801 | -99.5022
26.1839 | -99.4898
26.1876 | -99.4765
26.1913 | -99.4675
26.1951 | -99.4602
26.1988 | -99.4512
26.2026 | -99.4422
26.2063 | -99.4352
26.2101 | -99.4269
26.2138 | -99.4183
26.2175 | -99.4111
26.2213 | -99.4044
26.225 | -99.3994
26.2288 | -99.3952
26.2325 | -99.3902
26.2363 | -99.3867
26.24 | -99.3833
26.2437 | -99.3809
26.2475 | -99.3803
26.2512 | -99.3802
26.255 | -99.382
26.2587 | -99.3843
26.2625 | -99.388
26.2662 | -99.3925
26.2699 | -99.3971
26.2737 | -99.4029
26.2774 | -99.4093
26.2812 | -99.4168
26.2849 | -99.4233
26.2887 | -99.4301
26.2924 | -99.438
26.2961 | -99.4457
26.2999 | -99.4539
26.3036 | -99.4648
26.3074 | -99.4775
26.3111 | -99.491
26.3149 | -99.5044
26.3186 | -99.5172
26.3223 | -99.5303
26.3261 | -99.5411
26.3298 | -99.5502
26.3336 | -99.558
26.3373 | -99.5631
26.3411 | -99.5694
26.3448 | -99.5782
26.3485 | -99.588
26.3523 | -99.5972
26.356 | -99.607
26.3598 | -99.6159
26.3635 | -99.626
26.3673 | -99.634
26.371 | -99.6387
26.3747 | -99.6431
26.3785 | -99.6491
26.3822 | -99.6563
26.386 | -99.664
26.3897 | -99.6709
26.3935 | -99.674
26.3972 | -99.6763
26.4009 | -99.678
26.4047 | -99.678
26.4084 | -99.6776
26.4122 | -99.6778
26.4159 | -99.6771
26.4196 | -99.676
26.4234 | -99.6721
26.4271 | -99.6667
26.4309 | -99.6608
26.4346 | -99.6539
26.4384 | -99.6468
26.4421 | -99.6381
26.4458 | -99.6288
26.4496 | -99.6197
26.4533 | -99.6103
26.4571 | -99.5965
26.4608 | -99.586
26.4646 | -99.5733
26.4683 | -99.5615
26.472 | -99.5499
26.4758 | -99.5358
26.4795 | -99.5227
26.4833 | -99.5122
26.487 | -99.5008
26.4908 | -99.4902
26.4945 | -99.4805
26.4982 | -99.4723
26.502 | -99.4667
26.5057 | -99.4598
26.5095 | -99.4529
26.5132 | -99.4468
26.517 | -99.4423
26.5207 | -99.4384
26.5244 | -99.4352
26.5282 | -99.4338
26.5319 | -99.4314
26.5357 | -99.4282
26.5394 | -99.4273
26.5432 | -99.4297
26.5469 | -99.4345
26.5506 | -99.4407
26.5544 | -99.4493
26.5581 | -99.4584
26.5619 | -99.4698
26.5656 | -99.4828
26.5694 | -99.4982
26.5731 | -99.5165
26.5768 | -99.5356
26.5806 | -99.5545
26.5843 | -99.5759
26.5881 | -99.5989
26.5918 | -99.6205
26.5956 | -99.6439
26.5993 | -99.6685
26.603 | -99.6934
26.6068 | -99.7198
26.6105 | -99.748
26.6143 | -99.777
26.618 | -99.8069
26.6218 | -99.8365
26.6255 | -99.866
26.6292 | -99.8966
26.633 | -99.9289
26.6367 | -99.9591
26.6405 | -99.9878
26.6442 | -100.0162
26.648 | -100.0433
26.6517 | -100.0711
26.6554 | -100.1014
26.6592 | -100.1323
26.6629 | -100.1625
26.6667 | -100.1908
26.6704 | -100.2206
26.6742 | -100.2506
26.6779 | -100.2768
26.6816 | -100.3013
26.6854 | -100.3236
26.6891 | -100.3439
26.6929 | -100.3652
26.6966 | -100.387
26.7004 | -100.4083
26.7041 | -100.4304
26.7078 | -100.4534
26.7116 | -100.4749
26.7153 | -100.4922
26.7191 | -100.5058
26.7228 | -100.5153
26.7265 | -100.5225
26.7303 | -100.5284
26.734 | -100.5307
26.7378 | -100.5309
26.7415 | -100.5312
26.7453 | -100.5307
26.749 | -100.5273
26.7527 | -100.5209
26.7565 | -100.5131
26.7602 | -100.5036
26.764 | -100.4941
26.7677 | -100.483
26.7715 | -100.468
26.7752 | -100.4494
26.7789 | -100.432
26.7827 | -100.4155
26.7864 | -100.3982
26.7902 | -100.3801
26.7939 | -100.3614
26.7977 | -100.3438
26.8014 | -100.3253
26.8051 | -100.306
26.8089 | -100.2855
26.8126 | -100.2637
26.8164 | -100.2414
26.8201 | -100.2205
26.8239 | -100.2007
26.8276 | -100.1785
26.8313 | -100.1548
26.8351 | -100.1292
26.8388 | -100.1053
26.8426 | -100.0842
26.8463 | -100.0619
26.8501 | -100.0392
26.8538 | -100.0188
26.8575 | -99.9965
26.8613 | -99.9727
26.865 | -99.9495
26.8688 | -99.9259
26.8725 | -99.9059
26.8763 | -99.8869
26.88 | -99.8691
26.8837 | -99.8533
26.8875 | -99.8369
26.8912 | -99.8205
26.895 | -99.8055
26.8987 | -99.7907
26.9025 | -99.7761
26.9062 | -99.7618
26.9099 | -99.7495
26.9137 | -99.7416
26.9174 | -99.7337
26.9212 | -99.7254
26.9249 | -99.7182
26.9287 | -99.7124
26.9324 | -99.7064
26.9361 | -99.7025
26.9399 | -99.6995
26.9436 | -99.6971
26.9474 | -99.6968
26.9511 | -99.6977
26.9549 | -99.6988
26.9586 | -99.7038
26.9623 | -99.7111
26.9661 | -99.7194
26.9698 | -99.7281
26.9736 | -99.7361
26.9773 | -99.744
26.9811 | -99.7547
26.9848 | -99.7664
26.9885 | -99.7787
26.9923 | -99.7903
26.996 | -99.8074
26.9998 | -99.8208
27.0035 | -99.8347
27.0073 | -99.8495
27.011 | -99.8642
27.0147 | -99.8786
27.0185 | -99.8946
27.0222 | -99.9118
27.026 | -99.9294
27.0297 | -99.9462
27.0334 | -99.9607
27.0372 | -99.9722
27.0409 | -99.9823
27.0447 | -99.9923
27.0484 | -100.0046
27.0522 | -100.0156
27.0559 | -100.0247
27.0596 | -100.0341
27.0634 | -100.0437
27.0671 | -100.0546
27.0709 | -100.0651
27.0746 | -100.0732
27.0784 | -100.0799
27.0821 | -100.0872
27.0858 | -100.0922
27.0896 | -100.0957
27.0933 | -100.1003
27.0971 | -100.1044
27.1008 | -100.1081
27.1046 | -100.1125
27.1083 | -100.1176
27.112 | -100.123
27.1158 | -100.1276
27.1195 | -100.1306
27.1233 | -100.1322
27.127 | -100.1331
27.1308 | -100.1318
27.1345 | -100.1294
27.1382 | -100.1252
27.142 | -100.1215
27.1457 | -100.1159
27.1495 | -100.1112
27.1532 | -100.1103
27.157 | -100.1093
27.1607 | -100.107
27.1644 | -100.1037
27.1682 | -100.0998
27.1719 | -100.098
27.1757 | -100.095
27.1794 | -100.0895
27.1832 | -100.0866
27.1869 | -100.0849
27.1906 | -100.081
27.1944 | -100.0774
27.1981 | -100.0761
27.2019 | -100.073
27.2056 | -100.0688
27.2094 | -100.0659
27.2131 | -100.062
27.2168 | -100.0598
27.2206 | -100.0578
27.2243 | -100.0554
27.2281 | -100.0551
27.2318 | -100.0543
27.2356 | -100.0531
27.2393 | -100.052
27.243 | -100.0485
27.2468 | -100.044
27.2505 | -100.0412
27.2543 | -100.0361
27.258 | -100.0311
27.2618 | -100.0244
27.2655 | -100.0179
27.2692 | -100.0147
27.273 | -100.0115
27.2767 | -100.0057
27.2805 | -99.9994
27.2842 | -99.9949
27.288 | -99.9901
27.2917 | -99.9826
27.2954 | -99.972
27.2992 | -99.963
27.3029 | -99.9538
27.3067 | -99.9418
27.3104 | -99.9282
27.3142 | -99.9122
27.3179 | -99.8933
27.3216 | -99.8771
27.3254 | -99.8603
27.3291 | -99.8409
27.3329 | -99.8195
27.3366 | -99.7988
27.3404 | -99.7777
27.3441 | -99.7579
27.3478 | -99.7351
27.3516 | -99.7125
27.3553 | -99.6926
27.3591 | -99.672
27.3628 | -99.6488
27.3666 | -99.6243
27.3703 | -99.5989
27.374 | -99.574
27.3778 | -99.5495
27.3815 | -99.5219
27.3853 | -99.4928
27.389 | -99.4654
27.3927 | -99.4402
27.3965 | -99.4132
27.4002 | -99.3861
27.404 | -99.3599
27.4077 | -99.3348
27.4115 | -99.3099
27.4152 | -99.282
27.4189 | -99.2555
27.4227 | -99.232
27.4264 | -99.21
27.4302 | -99.187
27.4339 | -99.1621
27.4377 | -99.1378
27.4414 | -99.1147
27.4451 | -99.0924
27.4489 | -99.0699
27.4526 | -99.0484
27.4564 | -99.0288
27.4601 | -99.0105
27.4639 | -98.9929
27.4676 | -98.975
27.4713 | -98.9585
27.4751 | -98.9432
27.4788 | -98.929
27.4826 | -98.9157
27.4863 | -98.9022
27.4901 | -98.8888
27.4938 | -98.8764
27.4975 | -98.8635
27.5013 | -98.8523
27.505 | -98.8427
27.5088 | -98.8345
27.5125 | -98.8279
27.5163 | -98.8237
27.52 | -98.8206
27.5237 | -98.8177
27.5275 | -98.8151
27.5312 | -98.8152
27.535 | -98.8209
27.5387 | -98.8192
27.5425 | -98.8152
27.5462 | -98.8131
27.5499 | -98.8122
27.5537 | -98.8122
27.5574 | -98.813
27.5612 | -98.8139
27.5649 | -98.8169
27.5687 | -98.821
27.5724 | -98.8248
27.5761 | -98.8288
27.5799 | -98.8353
27.5836 | -98.8404
27.5874 | -98.8438
27.5911 | -98.8502
27.5949 | -98.8572
27.5986 | -98.8647
27.6023 | -98.8748
27.6061 | -98.8845
27.6098 | -98.8935
27.6136 | -98.9025
27.6173 | -98.9094
27.6211 | -98.9176
27.6248 | -98.9266
27.6285 | -98.9363
27.6323 | -98.9473
27.636 | -98.9599
27.6398 | -98.9707
27.6435 | -98.9802
27.6473 | -98.9876
27.651 | -98.9958
27.6547 | -99.0043
27.6585 | -99.0106
27.6622 | -99.0172
27.666 | -99.0249
27.6697 | -99.0315
27.6734 | -99.0356
27.6772 | -99.0383
27.6809 | -99.0421
27.6847 | -99.0471
27.6884 | -99.0498
27.6922 | -99.051
27.6959 | -99.0517
27.6996 | -99.0545
27.7034 | -99.0606
27.7071 | -99.0664
27.7109 | -99.0681
27.7146 | -99.0689
27.7184 | -99.0703
27.7221 | -99.0728
27.7258 | -99.074
27.7296 | -99.0748
27.7333 | -99.0737
27.7371 | -99.0748
27.7408 | -99.0747
27.7446 | -99.0712
27.7483 | -99.0688
27.752 | -99.0688
27.7558 | -99.0672
27.7595 | -99.0662
27.7633 | -99.0654
27.767 | -99.0617
27.7708 | -99.0597
27.7745 | -99.06
27.7782 | -99.0558
27.782 | -99.0511
27.7857 | -99.0498
27.7895 | -99.0474
27.7932 | -99.0432
27.797 | -99.0384
27.8007 | -99.0337
27.8044 | -99.0326
27.8082 | -99.0302
27.8119 | -99.0248
27.8157 | -99.0178
27.8194 | -99.012
27.8232 | -99.0071
27.8269 | -99.0017
27.8306 | -98.9958
27.8344 | -98.99
27.8381 | -98.9842
27.8419 | -98.978
27.8456 | -98.9692
27.8494 | -98.9613
27.8531 | -98.9531
27.8568 | -98.943
27.8606 | -98.9317
27.8643 | -98.9204
27.8681 | -98.9074
27.8718 | -98.8973
27.8756 | -98.8888
27.8793 | -98.8781
27.883 | -98.8677
27.8868 | -98.8596
27.8905 | -98.8524
27.8943 | -98.8434
27.898 | -98.833
27.9018 | -98.8211
27.9055 | -98.8099
27.9092 | -98.7972
27.913 | -98.7848
27.9167 | -98.7731
27.9205 | -98.7588
27.9242 | -98.7452
27.928 | -98.7296
27.9317 | -98.7104
27.9354 | -98.6972
27.9392 | -98.6855
27.9429 | -98.6725
27.9467 | -98.6576
27.9504 | -98.6425
27.9542 | -98.6255
27.9579 | -98.609
27.9616 | -98.5955
27.9654 | -98.5815
27.9691 | -98.5669
27.9729 | -98.5514
27.9766 | -98.536
27.9804 | -98.5207
27.9841 | -98.5048
27.9878 | -98.489
27.9916 | -98.4732
27.9953 | -98.4578
27.9991 | -98.4417
28.0028 | -98.4256