## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 268 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 271
Hood | 02-92B/P | 198
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1003.7 | hPa |  | PRC PRO 01
Humidity | 57 | %/Hr
Internal temperature of calibration bench | 24.11 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1011 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.6 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.07 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.96 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3113 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 02/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 20/07/2018
Technical Manager | Anthony HUE | Approval |  | 24/07/2018
CEO | Robert GLEMOT | Approval |  | 24/07/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.428 | 57.4318
0.012209 | -36.0245 | 38.2665
0.018313 | -34.9753 | 28.0932
0.024417 | -34.5106 | 21.8255
0.030521 | -34.3981 | 18.0751
0.036626 | -34.2571 | 15.0369
0.04273 | -34.1635 | 12.69
0.048834 | -34.1181 | 11.1602
0.054938 | -34.0562 | 10.0179
0.061043 | -33.9982 | 8.9435
0.067147 | -33.9879 | 8.1228
0.073251 | -33.9884 | 7.3665
0.079355 | -34.0054 | 6.7075
0.08546 | -34.0182 | 6.1175
0.091564 | -34.0134 | 5.6786
0.097668 | -33.9961 | 5.4226
0.10377 | -33.9755 | 5.1725
0.10988 | -33.9596 | 4.8959
0.11598 | -33.9514 | 4.6067
0.12209 | -33.9537 | 4.3269
0.12819 | -33.9639 | 4.0483
0.13429 | -33.9749 | 3.7673
0.1404 | -33.9781 | 3.4986
0.1465 | -33.9799 | 3.2482
0.15261 | -33.9772 | 3.033
0.15871 | -33.9708 | 2.8715
0.16482 | -33.963 | 2.7655
0.17092 | -33.961 | 2.7207
0.17702 | -33.9592 | 2.6452
0.18313 | -33.9587 | 2.5648
0.18923 | -33.9588 | 2.4558
0.19534 | -33.958 | 2.3303
0.20144 | -33.9572 | 2.1908
0.20754 | -33.9546 | 2.0725
0.21365 | -33.9538 | 1.9602
0.21975 | -33.9534 | 1.8642
0.22586 | -33.9543 | 1.7788
0.23196 | -33.9566 | 1.7077
0.23807 | -33.9604 | 1.6458
0.24417 | -33.9644 | 1.5874
0.25027 | -33.9682 | 1.5256
0.25638 | -33.9706 | 1.4641
0.26248 | -33.9725 | 1.3981
0.26859 | -33.9736 | 1.3332
0.27469 | -33.9741 | 1.2631
0.2808 | -33.9752 | 1.1916
0.2869 | -33.9762 | 1.1225
0.293 | -33.9772 | 1.0573
0.29911 | -33.9782 | 0.99873
0.30521 | -33.9792 | 0.94135
0.31132 | -33.98 | 0.9027
0.31742 | -33.9812 | 0.86012
0.32353 | -33.9816 | 0.8248
0.32963 | -33.9825 | 0.7949
0.33573 | -33.9836 | 0.77295
0.34184 | -33.9849 | 0.75021
0.34794 | -33.987 | 0.72556
0.35405 | -33.9891 | 0.70059
0.36015 | -33.9915 | 0.67456
0.36626 | -33.9935 | 0.64663
0.37236 | -33.9956 | 0.61031
0.37846 | -33.9977 | 0.56698
0.38457 | -33.9987 | 0.51182
0.39067 | -33.9998 | 0.46636
0.39678 | -34.0006 | 0.41603
0.40288 | -34.0011 | 0.36371
0.40899 | -34.0012 | 0.31336
0.41509 | -34.0007 | 0.264
0.42119 | -34.0008 | 0.22173
0.4273 | -34.0008 | 0.17673
0.4334 | -34.0008 | 0.13193
0.43951 | -34.0011 | 0.089654
0.44561 | -34.002 | 0.050778
0.45172 | -34.0034 | 0.019568
0.45782 | -34.0042 | -0.016743
0.46392 | -34.0053 | -0.052174
0.47003 | -34.0063 | -0.083942
0.47613 | -34.0071 | -0.1173
0.48224 | -34.0076 | -0.15128
0.48834 | -34.0067 | -0.18502
0.49445 | -34.0065 | -0.21924
0.50055 | -34.006 | -0.25058
0.50665 | -34.0054 | -0.28385
0.51276 | -34.0045 | -0.31684
0.51886 | -34.0035 | -0.34968
0.52497 | -34.0025 | -0.37328
0.53107 | -34.0018 | -0.40194
0.53717 | -34.0012 | -0.42923
0.54328 | -34.0005 | -0.45168
0.54938 | -34.0004 | -0.47362
0.55549 | -34.0005 | -0.49414
0.56159 | -34.0016 | -0.50604
0.5677 | -34.002 | -0.52069
0.5738 | -34.003 | -0.53706
0.5799 | -34.0038 | -0.55214
0.58601 | -34.0049 | -0.56272
0.59211 | -34.0065 | -0.57012
0.59822 | -34.0078 | -0.58031
0.60432 | -34.0093 | -0.59296
0.61043 | -34.0105 | -0.60178
0.61653 | -34.0117 | -0.61291
0.62263 | -34.0132 | -0.62322
0.62874 | -34.0138 | -0.63488
0.63484 | -34.0147 | -0.6451
0.64095 | -34.0152 | -0.65621
0.64705 | -34.0156 | -0.66698
0.65316 | -34.0158 | -0.6802
0.65926 | -34.0159 | -0.69655
0.66536 | -34.0149 | -0.71426
0.67147 | -34.0146 | -0.72801
0.67757 | -34.0143 | -0.74281
0.68368 | -34.0137 | -0.76188
0.68978 | -34.0131 | -0.77921
0.69589 | -34.0122 | -0.7982
0.70199 | -34.0114 | -0.8205
0.70809 | -34.0107 | -0.83953
0.7142 | -34.0101 | -0.85979
0.7203 | -34.0092 | -0.87881
0.72641 | -34.0086 | -0.90088
0.73251 | -34.0081 | -0.92056
0.73862 | -34.0084 | -0.94185
0.74472 | -34.0081 | -0.96161
0.75082 | -34.0079 | -0.98331
0.75693 | -34.0083 | -1.0047
0.76303 | -34.0089 | -1.0262
0.76914 | -34.0106 | -1.0464
0.77524 | -34.0116 | -1.0657
0.78135 | -34.0128 | -1.0856
0.78745 | -34.0141 | -1.1052
0.79355 | -34.0153 | -1.1248
0.79966 | -34.017 | -1.144
0.80576 | -34.0192 | -1.1634
0.81187 | -34.0207 | -1.1847
0.81797 | -34.0226 | -1.2056
0.82408 | -34.0242 | -1.2259
0.83018 | -34.0258 | -1.2462
0.83628 | -34.0276 | -1.2645
0.84239 | -34.0288 | -1.2847
0.84849 | -34.0301 | -1.3061
0.8546 | -34.0315 | -1.3256
0.8607 | -34.0327 | -1.346
0.86681 | -34.0337 | -1.3647
0.87291 | -34.0347 | -1.3833
0.87901 | -34.0351 | -1.4011
0.88512 | -34.0358 | -1.4191
0.89122 | -34.0365 | -1.4371
0.89733 | -34.0369 | -1.4527
0.90343 | -34.0373 | -1.47
0.90953 | -34.0375 | -1.4842
0.91564 | -34.0375 | -1.4999
0.92174 | -34.0377 | -1.5148
0.92785 | -34.0378 | -1.5294
0.93395 | -34.0379 | -1.5431
0.94006 | -34.0381 | -1.556
0.94616 | -34.0379 | -1.5679
0.95226 | -34.038 | -1.5805
0.95837 | -34.0381 | -1.5943
0.96447 | -34.0383 | -1.6071
0.97058 | -34.0381 | -1.6203
0.97668 | -34.038 | -1.6336
0.98279 | -34.0377 | -1.6463
0.98889 | -34.0376 | -1.6571
0.99499 | -34.0376 | -1.6704
1.0011 | -34.0375 | -1.6813
1.0072 | -34.0372 | -1.693
1.0133 | -34.0369 | -1.7039
1.0194 | -34.0365 | -1.7169
1.0255 | -34.0362 | -1.7306
1.0316 | -34.0357 | -1.7426
1.0377 | -34.0354 | -1.7563
1.0438 | -34.0349 | -1.7691
1.0499 | -34.0342 | -1.7844
1.056 | -34.0336 | -1.7987
1.0621 | -34.0331 | -1.8129
1.0682 | -34.0327 | -1.8255
1.0743 | -34.0321 | -1.8409
1.0805 | -34.0317 | -1.8553
1.0866 | -34.0309 | -1.8705
1.0927 | -34.0301 | -1.8851
1.0988 | -34.0296 | -1.9
1.1049 | -34.0289 | -1.9142
1.111 | -34.0282 | -1.9296
1.1171 | -34.0275 | -1.9442
1.1232 | -34.0271 | -1.9594
1.1293 | -34.0268 | -1.9735
1.1354 | -34.0266 | -1.9885
1.1415 | -34.0262 | -2.004
1.1476 | -34.0257 | -2.0195
1.1537 | -34.0253 | -2.0342
1.1598 | -34.0254 | -2.049
1.1659 | -34.0252 | -2.0624
1.172 | -34.025 | -2.0764
1.1781 | -34.0248 | -2.0904
1.1842 | -34.0247 | -2.1046
1.1903 | -34.025 | -2.1158
1.1964 | -34.0249 | -2.1293
1.2025 | -34.025 | -2.1432
1.2086 | -34.0253 | -2.1559
1.2147 | -34.0254 | -2.1685
1.2209 | -34.0257 | -2.1799
1.227 | -34.0264 | -2.1909
1.2331 | -34.0268 | -2.203
1.2392 | -34.0271 | -2.215
1.2453 | -34.0275 | -2.2262
1.2514 | -34.0279 | -2.2364
1.2575 | -34.0284 | -2.2474
1.2636 | -34.0292 | -2.2565
1.2697 | -34.0298 | -2.266
1.2758 | -34.0304 | -2.2772
1.2819 | -34.0311 | -2.2873
1.288 | -34.0317 | -2.2982
1.2941 | -34.0325 | -2.3069
1.3002 | -34.0331 | -2.3152
1.3063 | -34.0336 | -2.3244
1.3124 | -34.0343 | -2.3336
1.3185 | -34.0349 | -2.3428
1.3246 | -34.0356 | -2.352
1.3307 | -34.0364 | -2.3607
1.3368 | -34.0371 | -2.3695
1.3429 | -34.0377 | -2.3785
1.349 | -34.0382 | -2.3879
1.3551 | -34.0389 | -2.3973
1.3613 | -34.0394 | -2.4057
1.3674 | -34.04 | -2.4157
1.3735 | -34.0405 | -2.4255
1.3796 | -34.0409 | -2.4346
1.3857 | -34.0412 | -2.4439
1.3918 | -34.0417 | -2.4542
1.3979 | -34.0423 | -2.4634
1.404 | -34.0428 | -2.4746
1.4101 | -34.0432 | -2.485
1.4162 | -34.0436 | -2.4955
1.4223 | -34.0441 | -2.5068
1.4284 | -34.0445 | -2.5168
1.4345 | -34.0448 | -2.5283
1.4406 | -34.0451 | -2.5416
1.4467 | -34.0455 | -2.5531
1.4528 | -34.046 | -2.5644
1.4589 | -34.0465 | -2.5761
1.465 | -34.0467 | -2.587
1.4711 | -34.0472 | -2.5983
1.4772 | -34.0477 | -2.6114
1.4833 | -34.0482 | -2.6243
1.4894 | -34.0487 | -2.6363
1.4955 | -34.0493 | -2.6482
1.5016 | -34.0498 | -2.6606
1.5078 | -34.0502 | -2.6744
1.5139 | -34.0506 | -2.6873
1.52 | -34.051 | -2.6998
1.5261 | -34.0514 | -2.7124
1.5322 | -34.0517 | -2.7262
1.5383 | -34.0521 | -2.7387
1.5444 | -34.0526 | -2.7513
1.5505 | -34.0531 | -2.7646
1.5566 | -34.0536 | -2.7775
1.5627 | -34.0541 | -2.7911
1.5688 | -34.0547 | -2.8046
1.5749 | -34.055 | -2.8172
1.581 | -34.0553 | -2.8306
1.5871 | -34.0558 | -2.8442
1.5932 | -34.0561 | -2.8571
1.5993 | -34.0565 | -2.8712
1.6054 | -34.0569 | -2.8844
1.6115 | -34.0573 | -2.8983
1.6176 | -34.0576 | -2.911
1.6237 | -34.0579 | -2.9246
1.6298 | -34.0581 | -2.9372
1.6359 | -34.0585 | -2.95
1.642 | -34.0589 | -2.962
1.6482 | -34.0591 | -2.9751
1.6543 | -34.0595 | -2.9888
1.6604 | -34.0598 | -3.0008
1.6665 | -34.0599 | -3.0134
1.6726 | -34.0601 | -3.026
1.6787 | -34.0605 | -3.0387
1.6848 | -34.0604 | -3.0515
1.6909 | -34.0607 | -3.0643
1.697 | -34.0609 | -3.0766
1.7031 | -34.061 | -3.0893
1.7092 | -34.061 | -3.1016
1.7153 | -34.0612 | -3.1137
1.7214 | -34.061 | -3.1261
1.7275 | -34.061 | -3.1388
1.7336 | -34.0609 | -3.1518
1.7397 | -34.0611 | -3.1642
1.7458 | -34.0611 | -3.1766
1.7519 | -34.0611 | -3.1888
1.758 | -34.0612 | -3.201
1.7641 | -34.0614 | -3.2127
1.7702 | -34.0613 | -3.2257
1.7763 | -34.0613 | -3.2374
1.7824 | -34.0613 | -3.25
1.7885 | -34.0613 | -3.2625
1.7947 | -34.0615 | -3.2746
1.8008 | -34.0615 | -3.2861
1.8069 | -34.0615 | -3.2978
1.813 | -34.0616 | -3.3101
1.8191 | -34.0616 | -3.3229
1.8252 | -34.0617 | -3.3357
1.8313 | -34.0619 | -3.3479
1.8374 | -34.062 | -3.3599
1.8435 | -34.0621 | -3.3719
1.8496 | -34.0622 | -3.3849
1.8557 | -34.0624 | -3.3977
1.8618 | -34.0627 | -3.4098
1.8679 | -34.0628 | -3.4219
1.874 | -34.0629 | -3.4339
1.8801 | -34.0631 | -3.4462
1.8862 | -34.0634 | -3.4583
1.8923 | -34.0635 | -3.4704
1.8984 | -34.0637 | -3.4824
1.9045 | -34.0638 | -3.494
1.9106 | -34.0639 | -3.5061
1.9167 | -34.064 | -3.5183
1.9228 | -34.0641 | -3.5302
1.9289 | -34.0643 | -3.5419
1.9351 | -34.0645 | -3.5528
1.9412 | -34.0647 | -3.5641
1.9473 | -34.065 | -3.5761
1.9534 | -34.0652 | -3.5875
1.9595 | -34.0654 | -3.5993
1.9656 | -34.0657 | -3.6114
1.9717 | -34.0659 | -3.6231
1.9778 | -34.0663 | -3.6344
1.9839 | -34.0665 | -3.6458
1.99 | -34.0667 | -3.6569
1.9961 | -34.0669 | -3.6689
2.0022 | -34.0671 | -3.6802
2.0083 | -34.0672 | -3.6914
2.0144 | -34.0674 | -3.7028
2.0205 | -34.0678 | -3.7145
2.0266 | -34.0681 | -3.7258
2.0327 | -34.0683 | -3.7373
2.0388 | -34.0688 | -3.7488
2.0449 | -34.0689 | -3.7601
2.051 | -34.0692 | -3.7714
2.0571 | -34.0695 | -3.7827
2.0632 | -34.0698 | -3.7941
2.0693 | -34.0699 | -3.8061
2.0754 | -34.0702 | -3.8174
2.0816 | -34.0705 | -3.8295
2.0877 | -34.0707 | -3.8409
2.0938 | -34.0711 | -3.8529
2.0999 | -34.0712 | -3.8645
2.106 | -34.0714 | -3.8761
2.1121 | -34.0716 | -3.888
2.1182 | -34.0718 | -3.8987
2.1243 | -34.0721 | -3.9095
2.1304 | -34.0723 | -3.9206
2.1365 | -34.0726 | -3.9322
2.1426 | -34.0728 | -3.9441
2.1487 | -34.0731 | -3.9552
2.1548 | -34.0734 | -3.967
2.1609 | -34.0738 | -3.9783
2.167 | -34.0742 | -3.9895
2.1731 | -34.0745 | -4.0014
2.1792 | -34.0747 | -4.0124
2.1853 | -34.0751 | -4.0242
2.1914 | -34.0753 | -4.0355
2.1975 | -34.0757 | -4.0469
2.2036 | -34.0759 | -4.0589
2.2097 | -34.0761 | -4.0703
2.2158 | -34.0764 | -4.0809
2.222 | -34.0767 | -4.0913
2.2281 | -34.0769 | -4.1026
2.2342 | -34.0772 | -4.1139
2.2403 | -34.0774 | -4.1246
2.2464 | -34.0776 | -4.1354
2.2525 | -34.0778 | -4.1467
2.2586 | -34.0781 | -4.1573
2.2647 | -34.0784 | -4.1677
2.2708 | -34.0787 | -4.1789
2.2769 | -34.079 | -4.1896
2.283 | -34.0792 | -4.2007
2.2891 | -34.0794 | -4.2116
2.2952 | -34.0795 | -4.2222
2.3013 | -34.0797 | -4.2331
2.3074 | -34.08 | -4.2435
2.3135 | -34.0802 | -4.2549
2.3196 | -34.0803 | -4.2651
2.3257 | -34.0805 | -4.2757
2.3318 | -34.0807 | -4.2868
2.3379 | -34.0808 | -4.2975
2.344 | -34.081 | -4.3089
2.3501 | -34.0811 | -4.3196
2.3562 | -34.0813 | -4.3301
2.3623 | -34.0814 | -4.3405
2.3685 | -34.0815 | -4.3511
2.3746 | -34.0816 | -4.3623
2.3807 | -34.0818 | -4.3728
2.3868 | -34.082 | -4.3836
2.3929 | -34.0822 | -4.3942
2.399 | -34.0822 | -4.4047
2.4051 | -34.0825 | -4.4153
2.4112 | -34.0827 | -4.426
2.4173 | -34.0828 | -4.4366
2.4234 | -34.0829 | -4.4474
2.4295 | -34.0831 | -4.4583
2.4356 | -34.0832 | -4.4692
2.4417 | -34.0833 | -4.4804
2.4478 | -34.0834 | -4.4913
2.4539 | -34.0836 | -4.5016
2.46 | -34.0837 | -4.513
2.4661 | -34.0839 | -4.5241
2.4722 | -34.084 | -4.5344
2.4783 | -34.0841 | -4.5449
2.4844 | -34.0842 | -4.556
2.4905 | -34.0843 | -4.567
2.4966 | -34.0845 | -4.5777
2.5027 | -34.0846 | -4.5883
2.5089 | -34.0847 | -4.5992
2.515 | -34.0848 | -4.6101
2.5211 | -34.0849 | -4.6209
2.5272 | -34.085 | -4.6319
2.5333 | -34.0851 | -4.6429
2.5394 | -34.0853 | -4.6537
2.5455 | -34.0854 | -4.665
2.5516 | -34.0856 | -4.6755
2.5577 | -34.0858 | -4.6869
2.5638 | -34.0859 | -4.6987
2.5699 | -34.086 | -4.7101
2.576 | -34.086 | -4.7217
2.5821 | -34.0862 | -4.733
2.5882 | -34.0863 | -4.7441
2.5943 | -34.0864 | -4.755
2.6004 | -34.0867 | -4.766
2.6065 | -34.0868 | -4.7765
2.6126 | -34.0869 | -4.7871
2.6187 | -34.0871 | -4.7976
2.6248 | -34.0872 | -4.8085
2.6309 | -34.0874 | -4.82
2.637 | -34.0877 | -4.8316
2.6431 | -34.0879 | -4.8422
2.6492 | -34.088 | -4.8535
2.6554 | -34.0882 | -4.8644
2.6615 | -34.0883 | -4.8755
2.6676 | -34.0886 | -4.887
2.6737 | -34.0889 | -4.8982
2.6798 | -34.0891 | -4.9094
2.6859 | -34.0893 | -4.9207
2.692 | -34.0895 | -4.9316
2.6981 | -34.0897 | -4.9428
2.7042 | -34.09 | -4.9539
2.7103 | -34.0902 | -4.9647
2.7164 | -34.0905 | -4.9757
2.7225 | -34.0907 | -4.987
2.7286 | -34.091 | -4.9981
2.7347 | -34.0912 | -5.0087
2.7408 | -34.0915 | -5.0189
2.7469 | -34.0917 | -5.0302
2.753 | -34.092 | -5.0409
2.7591 | -34.0923 | -5.0522
2.7652 | -34.0926 | -5.0637
2.7713 | -34.0927 | -5.0746
2.7774 | -34.093 | -5.0855
2.7835 | -34.0932 | -5.0966
2.7896 | -34.0934 | -5.1071
2.7958 | -34.0937 | -5.1179
2.8019 | -34.094 | -5.1285
2.808 | -34.0943 | -5.1393
2.8141 | -34.0946 | -5.1506
2.8202 | -34.0949 | -5.1615
2.8263 | -34.0952 | -5.1725
2.8324 | -34.0956 | -5.1838
2.8385 | -34.0959 | -5.1947
2.8446 | -34.0963 | -5.2055
2.8507 | -34.0967 | -5.2164
2.8568 | -34.097 | -5.227
2.8629 | -34.0973 | -5.2376
2.869 | -34.0976 | -5.2487
2.8751 | -34.0979 | -5.2595
2.8812 | -34.0983 | -5.2709
2.8873 | -34.0986 | -5.2819
2.8934 | -34.0989 | -5.293
2.8995 | -34.0993 | -5.3041
2.9056 | -34.0995 | -5.3153
2.9117 | -34.0999 | -5.3264
2.9178 | -34.1002 | -5.3372
2.9239 | -34.1006 | -5.3487
2.93 | -34.101 | -5.3601
2.9361 | -34.1013 | -5.371
2.9423 | -34.1016 | -5.3817
2.9484 | -34.102 | -5.3928
2.9545 | -34.1024 | -5.4038
2.9606 | -34.1028 | -5.4152
2.9667 | -34.1031 | -5.4263
2.9728 | -34.1034 | -5.4376
2.9789 | -34.1039 | -5.4488
2.985 | -34.1043 | -5.46
2.9911 | -34.1046 | -5.4713
2.9972 | -34.1051 | -5.4825
3.0033 | -34.1055 | -5.494
3.0094 | -34.1058 | -5.5051
3.0155 | -34.1062 | -5.516
3.0216 | -34.1066 | -5.5273
3.0277 | -34.1069 | -5.539
3.0338 | -34.1073 | -5.5504
3.0399 | -34.1077 | -5.5617
3.046 | -34.108 | -5.5732
3.0521 | -34.1083 | -5.5843
3.0582 | -34.1086 | -5.5962
3.0643 | -34.1089 | -5.6076
3.0704 | -34.1092 | -5.6194
3.0765 | -34.1094 | -5.6305
3.0827 | -34.1097 | -5.6417
3.0888 | -34.11 | -5.6529
3.0949 | -34.1104 | -5.6641
3.101 | -34.1108 | -5.6756
3.1071 | -34.1111 | -5.6873
3.1132 | -34.1114 | -5.699
3.1193 | -34.1118 | -5.7104
3.1254 | -34.1121 | -5.7217
3.1315 | -34.1125 | -5.7336
3.1376 | -34.1127 | -5.7449
3.1437 | -34.1131 | -5.7565
3.1498 | -34.1134 | -5.7684
3.1559 | -34.1137 | -5.7798
3.162 | -34.1139 | -5.7918
3.1681 | -34.1142 | -5.8036
3.1742 | -34.1145 | -5.8151
3.1803 | -34.1148 | -5.8267
3.1864 | -34.1151 | -5.8379
3.1925 | -34.1154 | -5.8493
3.1986 | -34.1156 | -5.8612
3.2047 | -34.1159 | -5.8731
3.2108 | -34.1162 | -5.8842
3.2169 | -34.1165 | -5.8959
3.223 | -34.1167 | -5.9075
3.2292 | -34.1169 | -5.9193
3.2353 | -34.1171 | -5.9301
3.2414 | -34.1173 | -5.9414
3.2475 | -34.1175 | -5.9528
3.2536 | -34.1178 | -5.9642
3.2597 | -34.1181 | -5.9751
3.2658 | -34.1183 | -5.986
3.2719 | -34.1185 | -5.997
3.278 | -34.1187 | -6.0086
3.2841 | -34.119 | -6.0196
3.2902 | -34.1192 | -6.031
3.2963 | -34.1194 | -6.0424
3.3024 | -34.1195 | -6.0535
3.3085 | -34.1198 | -6.0649
3.3146 | -34.12 | -6.0761
3.3207 | -34.1202 | -6.0875
3.3268 | -34.1204 | -6.0989
3.3329 | -34.1206 | -6.1096
3.339 | -34.1208 | -6.1204
3.3451 | -34.121 | -6.1316
3.3512 | -34.1211 | -6.1423
3.3573 | -34.1214 | -6.1535
3.3634 | -34.1215 | -6.1646
3.3696 | -34.1217 | -6.1759
3.3757 | -34.1219 | -6.1866
3.3818 | -34.1221 | -6.1975
3.3879 | -34.1223 | -6.2086
3.394 | -34.1224 | -6.22
3.4001 | -34.1226 | -6.2308
3.4062 | -34.1227 | -6.2417
3.4123 | -34.1229 | -6.2522
3.4184 | -34.1231 | -6.2635
3.4245 | -34.1233 | -6.2743
3.4306 | -34.1234 | -6.285
3.4367 | -34.1236 | -6.2956
3.4428 | -34.1238 | -6.306
3.4489 | -34.1239 | -6.317
3.455 | -34.124 | -6.3278
3.4611 | -34.1242 | -6.3391
3.4672 | -34.1244 | -6.3502
3.4733 | -34.1246 | -6.3612
3.4794 | -34.1247 | -6.3721
3.4855 | -34.1249 | -6.3826
3.4916 | -34.1251 | -6.393
3.4977 | -34.1252 | -6.4042
3.5038 | -34.1254 | -6.4151
3.5099 | -34.1255 | -6.4255
3.5161 | -34.1257 | -6.4359
3.5222 | -34.1259 | -6.4464
3.5283 | -34.126 | -6.4572
3.5344 | -34.1261 | -6.468
3.5405 | -34.1263 | -6.4786
3.5466 | -34.1265 | -6.4895
3.5527 | -34.1266 | -6.5004
3.5588 | -34.1267 | -6.5111
3.5649 | -34.1268 | -6.5219
3.571 | -34.1269 | -6.5325
3.5771 | -34.1271 | -6.5436
3.5832 | -34.1272 | -6.5536
3.5893 | -34.1274 | -6.5643
3.5954 | -34.1276 | -6.5751
3.6015 | -34.1278 | -6.586
3.6076 | -34.128 | -6.5966
3.6137 | -34.1282 | -6.6068
3.6198 | -34.1284 | -6.6172
3.6259 | -34.1286 | -6.6282
3.632 | -34.1288 | -6.6389
3.6381 | -34.129 | -6.6497
3.6442 | -34.1291 | -6.6603
3.6503 | -34.1293 | -6.6706
3.6565 | -34.1296 | -6.6814
3.6626 | -34.1297 | -6.692
3.6687 | -34.13 | -6.7028
3.6748 | -34.1302 | -6.7133
3.6809 | -34.1305 | -6.724
3.687 | -34.1306 | -6.735
3.6931 | -34.1309 | -6.7453
3.6992 | -34.1311 | -6.756
3.7053 | -34.1313 | -6.7666
3.7114 | -34.1315 | -6.7774
3.7175 | -34.1317 | -6.7881
3.7236 | -34.132 | -6.7983
3.7297 | -34.1322 | -6.8091
3.7358 | -34.1324 | -6.8201
3.7419 | -34.1326 | -6.8309
3.748 | -34.1328 | -6.8415
3.7541 | -34.133 | -6.8522
3.7602 | -34.1333 | -6.8628
3.7663 | -34.1335 | -6.8733
3.7724 | -34.1338 | -6.8841
3.7785 | -34.1341 | -6.8946
3.7846 | -34.1342 | -6.9051
3.7907 | -34.1345 | -6.916
3.7969 | -34.1348 | -6.9268
3.803 | -34.135 | -6.9371
3.8091 | -34.1353 | -6.9477
3.8152 | -34.1356 | -6.9584
3.8213 | -34.1358 | -6.9692
3.8274 | -34.136 | -6.9796
3.8335 | -34.1364 | -6.9904
3.8396 | -34.1367 | -7.0012
3.8457 | -34.137 | -7.012
3.8518 | -34.1373 | -7.0225
3.8579 | -34.1376 | -7.0331
3.864 | -34.1379 | -7.0439
3.8701 | -34.1382 | -7.055
3.8762 | -34.1385 | -7.0655
3.8823 | -34.1387 | -7.0763
3.8884 | -34.139 | -7.087
3.8945 | -34.1394 | -7.0979
3.9006 | -34.1396 | -7.1089
3.9067 | -34.1399 | -7.1201
3.9128 | -34.1402 | -7.1308
3.9189 | -34.1404 | -7.1414
3.925 | -34.1407 | -7.1523
3.9311 | -34.1409 | -7.1626
3.9372 | -34.1412 | -7.1735
3.9434 | -34.1415 | -7.1842
3.9495 | -34.1419 | -7.195
3.9556 | -34.1422 | -7.2056
3.9617 | -34.1425 | -7.2159
3.9678 | -34.1428 | -7.2264
3.9739 | -34.1431 | -7.237
3.98 | -34.1433 | -7.2478
3.9861 | -34.1436 | -7.2585
3.9922 | -34.1438 | -7.2692
3.9983 | -34.1441 | -7.2795
4.0044 | -34.1444 | -7.2898
4.0105 | -34.1446 | -7.3005
4.0166 | -34.145 | -7.3112
4.0227 | -34.1452 | -7.3218
4.0288 | -34.1455 | -7.3329
4.0349 | -34.1457 | -7.3434
4.041 | -34.146 | -7.354
4.0471 | -34.1462 | -7.3648
4.0532 | -34.1465 | -7.3751
4.0593 | -34.1468 | -7.386
4.0654 | -34.1471 | -7.3964
4.0715 | -34.1473 | -7.4075
4.0776 | -34.1476 | -7.4181
4.0838 | -34.1479 | -7.4291
4.0899 | -34.1482 | -7.4396
4.096 | -34.1484 | -7.4502
4.1021 | -34.1487 | -7.4606
4.1082 | -34.1489 | -7.471
4.1143 | -34.1492 | -7.4813
4.1204 | -34.1495 | -7.4916
4.1265 | -34.1497 | -7.502
4.1326 | -34.15 | -7.5124
4.1387 | -34.1503 | -7.5229
4.1448 | -34.1505 | -7.5335
4.1509 | -34.1507 | -7.5442
4.157 | -34.151 | -7.5552
4.1631 | -34.1513 | -7.566
4.1692 | -34.1515 | -7.5765
4.1753 | -34.1518 | -7.587
4.1814 | -34.152 | -7.5973
4.1875 | -34.1523 | -7.6081
4.1936 | -34.1525 | -7.6187
4.1997 | -34.1527 | -7.629
4.2058 | -34.153 | -7.6398
4.2119 | -34.1532 | -7.6502
4.218 | -34.1534 | -7.6605
4.2241 | -34.1537 | -7.6708
4.2303 | -34.154 | -7.6812
4.2364 | -34.1542 | -7.6917
4.2425 | -34.1544 | -7.7023
4.2486 | -34.1546 | -7.7129
4.2547 | -34.1547 | -7.7233
4.2608 | -34.1549 | -7.7339
4.2669 | -34.1551 | -7.7442
4.273 | -34.1554 | -7.7548
4.2791 | -34.1556 | -7.7653
4.2852 | -34.1558 | -7.7759
4.2913 | -34.156 | -7.7862
4.2974 | -34.1562 | -7.797
4.3035 | -34.1565 | -7.8077
4.3096 | -34.1566 | -7.8184
4.3157 | -34.1569 | -7.8288
4.3218 | -34.1571 | -7.8392
4.3279 | -34.1573 | -7.8497
4.334 | -34.1575 | -7.8601
4.3401 | -34.1577 | -7.8702
4.3462 | -34.1579 | -7.8809
4.3523 | -34.1582 | -7.891
4.3584 | -34.1584 | -7.9011
4.3645 | -34.1586 | -7.9114
4.3707 | -34.1588 | -7.9216
4.3768 | -34.159 | -7.932
4.3829 | -34.1592 | -7.9428
4.389 | -34.1594 | -7.9531
4.3951 | -34.1596 | -7.9634
4.4012 | -34.1598 | -7.9741
4.4073 | -34.16 | -7.9843
4.4134 | -34.1602 | -7.9947
4.4195 | -34.1605 | -8.005
4.4256 | -34.1607 | -8.0152
4.4317 | -34.1609 | -8.0256
4.4378 | -34.1611 | -8.0357
4.4439 | -34.1613 | -8.0457
4.45 | -34.1615 | -8.0561
4.4561 | -34.1616 | -8.0664
4.4622 | -34.1618 | -8.0768
4.4683 | -34.162 | -8.0872
4.4744 | -34.1622 | -8.0978
4.4805 | -34.1624 | -8.1078
4.4866 | -34.1627 | -8.1181
4.4927 | -34.1629 | -8.1287
4.4988 | -34.1631 | -8.139
4.5049 | -34.1633 | -8.149
4.511 | -34.1636 | -8.159
4.5172 | -34.1638 | -8.1694
4.5233 | -34.1639 | -8.18
4.5294 | -34.1642 | -8.1902
4.5355 | -34.1644 | -8.2006
4.5416 | -34.1646 | -8.2108
4.5477 | -34.1648 | -8.2212
4.5538 | -34.165 | -8.2312
4.5599 | -34.1653 | -8.2414
4.566 | -34.1655 | -8.2517
4.5721 | -34.1657 | -8.2621
4.5782 | -34.166 | -8.2726
4.5843 | -34.1662 | -8.283
4.5904 | -34.1665 | -8.2934
4.5965 | -34.1667 | -8.3035
4.6026 | -34.1669 | -8.3142
4.6087 | -34.1671 | -8.3245
4.6148 | -34.1673 | -8.3352
4.6209 | -34.1676 | -8.3454
4.627 | -34.1678 | -8.3559
4.6331 | -34.168 | -8.3661
4.6392 | -34.1682 | -8.3765
4.6453 | -34.1685 | -8.3869
4.6514 | -34.1687 | -8.3973
4.6576 | -34.1689 | -8.408
4.6637 | -34.1691 | -8.418
4.6698 | -34.1693 | -8.4281
4.6759 | -34.1695 | -8.4389
4.682 | -34.1697 | -8.4492
4.6881 | -34.1699 | -8.4595
4.6942 | -34.1702 | -8.4698
4.7003 | -34.1704 | -8.4802
4.7064 | -34.1706 | -8.4904
4.7125 | -34.1709 | -8.501
4.7186 | -34.1711 | -8.5112
4.7247 | -34.1713 | -8.5216
4.7308 | -34.1715 | -8.5319
4.7369 | -34.1718 | -8.5423
4.743 | -34.172 | -8.553
4.7491 | -34.1722 | -8.5637
4.7552 | -34.1724 | -8.5742
4.7613 | -34.1726 | -8.5846
4.7674 | -34.1728 | -8.5948
4.7735 | -34.173 | -8.6056
4.7796 | -34.1734 | -8.6163
4.7857 | -34.1736 | -8.6267
4.7918 | -34.1738 | -8.6369
4.7979 | -34.1741 | -8.6471
4.8041 | -34.1743 | -8.6576
4.8102 | -34.1746 | -8.6677
4.8163 | -34.1748 | -8.6781
4.8224 | -34.175 | -8.6885
4.8285 | -34.1753 | -8.6987
4.8346 | -34.1756 | -8.7089
4.8407 | -34.1757 | -8.7194
4.8468 | -34.176 | -8.73
4.8529 | -34.1763 | -8.7407
4.859 | -34.1765 | -8.7515
4.8651 | -34.1767 | -8.7617
4.8712 | -34.177 | -8.7719
4.8773 | -34.1772 | -8.7824
4.8834 | -34.1774 | -8.7928
4.8895 | -34.1777 | -8.803
4.8956 | -34.1779 | -8.8135
4.9017 | -34.1781 | -8.8237
4.9078 | -34.1783 | -8.8345
4.9139 | -34.1786 | -8.8445
4.92 | -34.1789 | -8.8546
4.9261 | -34.1791 | -8.8647
4.9322 | -34.1794 | -8.8751
4.9383 | -34.1796 | -8.8855
4.9445 | -34.1799 | -8.8958
4.9506 | -34.1801 | -8.9061
4.9567 | -34.1804 | -8.9164
4.9628 | -34.1806 | -8.9267
4.9689 | -34.1808 | -8.9375
4.975 | -34.1811 | -8.9482
4.9811 | -34.1813 | -8.9583
4.9872 | -34.1816 | -8.9686
4.9933 | -34.1819 | -8.9792
4.9994 | -34.1822 | -8.9901
5.0055 | -34.1825 | -9.0005
5.0116 | -34.1827 | -9.0113
5.0177 | -34.1831 | -9.0219
5.0238 | -34.1833 | -9.0326
5.0299 | -34.1837 | -9.0433
5.036 | -34.1839 | -9.0537
5.0421 | -34.1842 | -9.064
5.0482 | -34.1844 | -9.0745
5.0543 | -34.1847 | -9.0847
5.0604 | -34.185 | -9.0952
5.0665 | -34.1853 | -9.1053
5.0726 | -34.1856 | -9.1161
5.0787 | -34.1859 | -9.1267
5.0848 | -34.1862 | -9.1373
5.091 | -34.1864 | -9.1476
5.0971 | -34.1867 | -9.1577
5.1032 | -34.187 | -9.1685
5.1093 | -34.1873 | -9.1789
5.1154 | -34.1875 | -9.1892
5.1215 | -34.1878 | -9.1995
5.1276 | -34.1881 | -9.2098
5.1337 | -34.1884 | -9.22
5.1398 | -34.1887 | -9.2302
5.1459 | -34.189 | -9.2409
5.152 | -34.1892 | -9.2515
5.1581 | -34.1895 | -9.2621
5.1642 | -34.1898 | -9.2726
5.1703 | -34.1901 | -9.283
5.1764 | -34.1904 | -9.2932
5.1825 | -34.1907 | -9.3035
5.1886 | -34.191 | -9.3136
5.1947 | -34.1912 | -9.3241
5.2008 | -34.1916 | -9.3346
5.2069 | -34.1918 | -9.3452
5.213 | -34.1921 | -9.3554
5.2191 | -34.1924 | -9.3659
5.2252 | -34.1927 | -9.3765
5.2314 | -34.193 | -9.3871
5.2375 | -34.1933 | -9.3974
5.2436 | -34.1935 | -9.4078
5.2497 | -34.1938 | -9.418
5.2558 | -34.1941 | -9.4288
5.2619 | -34.1944 | -9.4394
5.268 | -34.1947 | -9.45
5.2741 | -34.1951 | -9.4606
5.2802 | -34.1953 | -9.4714
5.2863 | -34.1957 | -9.4817
5.2924 | -34.196 | -9.4924
5.2985 | -34.1963 | -9.5029
5.3046 | -34.1966 | -9.5134
5.3107 | -34.1969 | -9.5241
5.3168 | -34.1972 | -9.5347
5.3229 | -34.1975 | -9.5452
5.329 | -34.1978 | -9.5557
5.3351 | -34.1981 | -9.566
5.3412 | -34.1984 | -9.5763
5.3473 | -34.1987 | -9.5869
5.3534 | -34.199 | -9.5975
5.3595 | -34.1993 | -9.608
5.3656 | -34.1997 | -9.6185
5.3717 | -34.2 | -9.6293
5.3779 | -34.2003 | -9.6403
5.384 | -34.2006 | -9.6507
5.3901 | -34.2009 | -9.6615
5.3962 | -34.2012 | -9.6718
5.4023 | -34.2015 | -9.6824
5.4084 | -34.2019 | -9.6928
5.4145 | -34.2022 | -9.7032
5.4206 | -34.2025 | -9.7138
5.4267 | -34.2028 | -9.7246
5.4328 | -34.2031 | -9.7351
5.4389 | -34.2034 | -9.7454
5.445 | -34.2037 | -9.7562
5.4511 | -34.2041 | -9.7666
5.4572 | -34.2044 | -9.7773
5.4633 | -34.2047 | -9.7878
5.4694 | -34.2049 | -9.7983
5.4755 | -34.2052 | -9.8085
5.4816 | -34.2055 | -9.8191
5.4877 | -34.2059 | -9.8294
5.4938 | -34.2062 | -9.8398
5.4999 | -34.2065 | -9.8502
5.506 | -34.2068 | -9.8607
5.5121 | -34.2071 | -9.8711
5.5183 | -34.2075 | -9.8814
5.5244 | -34.2077 | -9.8921
5.5305 | -34.2081 | -9.9025
5.5366 | -34.2084 | -9.9127
5.5427 | -34.2087 | -9.9234
5.5488 | -34.209 | -9.934
5.5549 | -34.2093 | -9.9446
5.561 | -34.2096 | -9.9553
5.5671 | -34.2099 | -9.9657
5.5732 | -34.2102 | -9.9764
5.5793 | -34.2106 | -9.9872
5.5854 | -34.2109 | -9.9976
5.5915 | -34.2112 | -10.0084
5.5976 | -34.2115 | -10.0189
5.6037 | -34.2118 | -10.0296
5.6098 | -34.2121 | -10.04
5.6159 | -34.2124 | -10.0503
5.622 | -34.2127 | -10.0608
5.6281 | -34.213 | -10.0714
5.6342 | -34.2134 | -10.0821
5.6403 | -34.2137 | -10.0927
5.6464 | -34.214 | -10.1034
5.6525 | -34.2143 | -10.114
5.6586 | -34.2146 | -10.1243
5.6648 | -34.2149 | -10.1349
5.6709 | -34.2153 | -10.1457
5.677 | -34.2156 | -10.1563
5.6831 | -34.2159 | -10.1668
5.6892 | -34.2162 | -10.1776
5.6953 | -34.2165 | -10.1881
5.7014 | -34.2168 | -10.1986
5.7075 | -34.2171 | -10.2093
5.7136 | -34.2174 | -10.2198
5.7197 | -34.2177 | -10.2302
5.7258 | -34.218 | -10.2406
5.7319 | -34.2183 | -10.2512
5.738 | -34.2186 | -10.2618
5.7441 | -34.2189 | -10.2722
5.7502 | -34.2193 | -10.2825
5.7563 | -34.2196 | -10.293
5.7624 | -34.22 | -10.3038
5.7685 | -34.2203 | -10.3145
5.7746 | -34.2206 | -10.3253
5.7807 | -34.2208 | -10.3359
5.7868 | -34.2211 | -10.3466
5.7929 | -34.2214 | -10.3571
5.799 | -34.2218 | -10.3675
5.8052 | -34.222 | -10.378
5.8113 | -34.2224 | -10.3885
5.8174 | -34.2226 | -10.3988
5.8235 | -34.223 | -10.4093
5.8296 | -34.2233 | -10.42
5.8357 | -34.2235 | -10.4303
5.8418 | -34.2238 | -10.441
5.8479 | -34.2241 | -10.4517
5.854 | -34.2244 | -10.4623
5.8601 | -34.2248 | -10.4728
5.8662 | -34.2251 | -10.4834
5.8723 | -34.2254 | -10.4938
5.8784 | -34.2256 | -10.5043
5.8845 | -34.2259 | -10.5147
5.8906 | -34.2262 | -10.5253
5.8967 | -34.2266 | -10.5356
5.9028 | -34.2269 | -10.546
5.9089 | -34.2273 | -10.5565
5.915 | -34.2276 | -10.5667
5.9211 | -34.2279 | -10.5772
5.9272 | -34.2282 | -10.5874
5.9333 | -34.2285 | -10.5978
5.9394 | -34.2288 | -10.6082
5.9455 | -34.2291 | -10.6185
5.9517 | -34.2295 | -10.6289
5.9578 | -34.2298 | -10.6395
5.9639 | -34.2301 | -10.6497
5.97 | -34.2305 | -10.6603
5.9761 | -34.2308 | -10.6706
5.9822 | -34.2311 | -10.6812
5.9883 | -34.2314 | -10.6918
5.9944 | -34.2317 | -10.7023
6.0005 | -34.2321 | -10.713
6.0066 | -34.2324 | -10.7232
6.0127 | -34.2327 | -10.7334
6.0188 | -34.233 | -10.744
6.0249 | -34.2333 | -10.7545
6.031 | -34.2336 | -10.7647
6.0371 | -34.2339 | -10.7751
6.0432 | -34.2342 | -10.7854
6.0493 | -34.2346 | -10.7957
6.0554 | -34.2349 | -10.8059
6.0615 | -34.2352 | -10.8164
6.0676 | -34.2355 | -10.8271
6.0737 | -34.2359 | -10.8376
6.0798 | -34.2361 | -10.8481
6.0859 | -34.2365 | -10.8583
6.0921 | -34.2368 | -10.8686
6.0982 | -34.2372 | -10.8791
6.1043 | -34.2375 | -10.8894
6.1104 | -34.2378 | -10.8999
6.1165 | -34.2382 | -10.9101
6.1226 | -34.2385 | -10.9205
6.1287 | -34.2388 | -10.9308
6.1348 | -34.2391 | -10.9413
6.1409 | -34.2394 | -10.9515
6.147 | -34.2398 | -10.9619
6.1531 | -34.2401 | -10.9724
6.1592 | -34.2404 | -10.9826
6.1653 | -34.2406 | -10.9929
6.1714 | -34.2409 | -11.0035
6.1775 | -34.2413 | -11.0138
6.1836 | -34.2416 | -11.0241
6.1897 | -34.2419 | -11.0344
6.1958 | -34.2422 | -11.0449
6.2019 | -34.2425 | -11.0555
6.208 | -34.2429 | -11.066
6.2141 | -34.2432 | -11.0763
6.2202 | -34.2435 | -11.0869
6.2263 | -34.2439 | -11.0973
6.2325 | -34.2441 | -11.1073
6.2386 | -34.2445 | -11.1173
6.2447 | -34.2448 | -11.1278
6.2508 | -34.2451 | -11.1381
6.2569 | -34.2454 | -11.1484
6.263 | -34.2457 | -11.1585
6.2691 | -34.246 | -11.1686
6.2752 | -34.2463 | -11.1789
6.2813 | -34.2467 | -11.1893
6.2874 | -34.247 | -11.1997
6.2935 | -34.2473 | -11.21
6.2996 | -34.2476 | -11.2201
6.3057 | -34.2479 | -11.2306
6.3118 | -34.2482 | -11.2411
6.3179 | -34.2485 | -11.2516
6.324 | -34.2488 | -11.2619
6.3301 | -34.2491 | -11.2722
6.3362 | -34.2495 | -11.2826
6.3423 | -34.2498 | -11.293
6.3484 | -34.2501 | -11.3032
6.3545 | -34.2504 | -11.3133
6.3606 | -34.2507 | -11.3231
6.3667 | -34.2509 | -11.333
6.3728 | -34.2513 | -11.3435
6.379 | -34.2516 | -11.3535
6.3851 | -34.2519 | -11.3633
6.3912 | -34.2522 | -11.3738
6.3973 | -34.2525 | -11.3841
6.4034 | -34.2528 | -11.3943
6.4095 | -34.2531 | -11.4047
6.4156 | -34.2534 | -11.4153
6.4217 | -34.2537 | -11.4257
6.4278 | -34.254 | -11.4359
6.4339 | -34.2544 | -11.4465
6.44 | -34.2547 | -11.4568
6.4461 | -34.255 | -11.4673
6.4522 | -34.2553 | -11.4772
6.4583 | -34.2555 | -11.4878
6.4644 | -34.2558 | -11.498
6.4705 | -34.2561 | -11.508
6.4766 | -34.2564 | -11.5183
6.4827 | -34.2567 | -11.5288
6.4888 | -34.257 | -11.5391
6.4949 | -34.2573 | -11.5493
6.501 | -34.2576 | -11.5598
6.5071 | -34.2579 | -11.57
6.5132 | -34.2582 | -11.5803
6.5194 | -34.2585 | -11.5907
6.5255 | -34.2588 | -11.6007
6.5316 | -34.2591 | -11.6111
6.5377 | -34.2593 | -11.6214
6.5438 | -34.2596 | -11.6317
6.5499 | -34.2599 | -11.6418
6.556 | -34.2602 | -11.652
6.5621 | -34.2605 | -11.6625
6.5682 | -34.2608 | -11.6728
6.5743 | -34.2611 | -11.683
6.5804 | -34.2614 | -11.6935
6.5865 | -34.2617 | -11.7038
6.5926 | -34.262 | -11.7139
6.5987 | -34.2623 | -11.724
6.6048 | -34.2626 | -11.7342
6.6109 | -34.2629 | -11.7445
6.617 | -34.2632 | -11.7549
6.6231 | -34.2635 | -11.7651
6.6292 | -34.2638 | -11.7754
6.6353 | -34.264 | -11.7859
6.6414 | -34.2644 | -11.7961
6.6475 | -34.2646 | -11.8064
6.6536 | -34.2649 | -11.8166
6.6597 | -34.2652 | -11.8265
6.6659 | -34.2655 | -11.8367
6.672 | -34.2658 | -11.8469
6.6781 | -34.266 | -11.8572
6.6842 | -34.2663 | -11.8673
6.6903 | -34.2667 | -11.8773
6.6964 | -34.267 | -11.8876
6.7025 | -34.2672 | -11.8978
6.7086 | -34.2675 | -11.9079
6.7147 | -34.2678 | -11.9181
6.7208 | -34.268 | -11.9282
6.7269 | -34.2683 | -11.9385
6.733 | -34.2686 | -11.9485
6.7391 | -34.2689 | -11.9589
6.7452 | -34.2692 | -11.9688
6.7513 | -34.2695 | -11.9788
6.7574 | -34.2698 | -11.9893
6.7635 | -34.2701 | -11.9992
6.7696 | -34.2704 | -12.0091
6.7757 | -34.2707 | -12.0193
6.7818 | -34.271 | -12.0297
6.7879 | -34.2713 | -12.0398
6.794 | -34.2716 | -12.0498
6.8001 | -34.2718 | -12.0598
6.8063 | -34.2721 | -12.0701
6.8124 | -34.2723 | -12.0805
6.8185 | -34.2726 | -12.091
6.8246 | -34.2728 | -12.1008
6.8307 | -34.2731 | -12.1108
6.8368 | -34.2734 | -12.1213
6.8429 | -34.2737 | -12.1313
6.849 | -34.274 | -12.1418
6.8551 | -34.2743 | -12.1515
6.8612 | -34.2746 | -12.1618
6.8673 | -34.2748 | -12.1722
6.8734 | -34.2751 | -12.1824
6.8795 | -34.2754 | -12.1927
6.8856 | -34.2757 | -12.2027
6.8917 | -34.276 | -12.2127
6.8978 | -34.2762 | -12.2225
6.9039 | -34.2765 | -12.2322
6.91 | -34.2769 | -12.2426
6.9161 | -34.2772 | -12.2524
6.9222 | -34.2774 | -12.2626
6.9283 | -34.2777 | -12.2722
6.9344 | -34.278 | -12.2823
6.9405 | -34.2783 | -12.2925
6.9466 | -34.2786 | -12.3026
6.9528 | -34.2789 | -12.3128
6.9589 | -34.2792 | -12.3231
6.965 | -34.2794 | -12.3331
6.9711 | -34.2797 | -12.3436
6.9772 | -34.28 | -12.3537
6.9833 | -34.2803 | -12.3639
6.9894 | -34.2806 | -12.3743
6.9955 | -34.2809 | -12.3842
7.0016 | -34.2812 | -12.3942
7.0077 | -34.2815 | -12.4042
7.0138 | -34.2818 | -12.4143
7.0199 | -34.282 | -12.4242
7.026 | -34.2823 | -12.4343
7.0321 | -34.2826 | -12.4441
7.0382 | -34.2829 | -12.4538
7.0443 | -34.2832 | -12.4641
7.0504 | -34.2835 | -12.4741
7.0565 | -34.2837 | -12.4842
7.0626 | -34.284 | -12.494
7.0687 | -34.2843 | -12.5035
7.0748 | -34.2846 | -12.5132
7.0809 | -34.2849 | -12.5232
7.087 | -34.2852 | -12.5333
7.0932 | -34.2854 | -12.5432
7.0993 | -34.2857 | -12.5531
7.1054 | -34.286 | -12.563
7.1115 | -34.2863 | -12.573
7.1176 | -34.2865 | -12.5827
7.1237 | -34.2868 | -12.5926
7.1298 | -34.2871 | -12.6027
7.1359 | -34.2873 | -12.6128
7.142 | -34.2876 | -12.6224
7.1481 | -34.2878 | -12.6324
7.1542 | -34.2881 | -12.6424
7.1603 | -34.2884 | -12.6523
7.1664 | -34.2887 | -12.662
7.1725 | -34.289 | -12.6719
7.1786 | -34.2893 | -12.6819
7.1847 | -34.2896 | -12.6917
7.1908 | -34.2898 | -12.7016
7.1969 | -34.2901 | -12.7115
7.203 | -34.2904 | -12.7212
7.2091 | -34.2907 | -12.7309
7.2152 | -34.291 | -12.7406
7.2213 | -34.2913 | -12.75
7.2274 | -34.2916 | -12.7601
7.2335 | -34.2918 | -12.7698
7.2397 | -34.2921 | -12.7795
7.2458 | -34.2924 | -12.789
7.2519 | -34.2926 | -12.799
7.258 | -34.2929 | -12.8089
7.2641 | -34.2932 | -12.8186
7.2702 | -34.2935 | -12.8283
7.2763 | -34.2938 | -12.8384
7.2824 | -34.2941 | -12.8484
7.2885 | -34.2943 | -12.858
7.2946 | -34.2946 | -12.8677
7.3007 | -34.2949 | -12.8778
7.3068 | -34.2951 | -12.8881
7.3129 | -34.2955 | -12.8976
7.319 | -34.2958 | -12.9075
7.3251 | -34.2961 | -12.9172
7.3312 | -34.2963 | -12.9268
7.3373 | -34.2966 | -12.9366
7.3434 | -34.2969 | -12.9465
7.3495 | -34.2973 | -12.956
7.3556 | -34.2975 | -12.966
7.3617 | -34.2978 | -12.9763
7.3678 | -34.2981 | -12.9859
7.3739 | -34.2984 | -12.9959
7.3801 | -34.2987 | -13.006
7.3862 | -34.299 | -13.0157
7.3923 | -34.2994 | -13.0255
7.3984 | -34.2997 | -13.0353
7.4045 | -34.3 | -13.0452
7.4106 | -34.3003 | -13.0551
7.4167 | -34.3005 | -13.0648
7.4228 | -34.3008 | -13.0752
7.4289 | -34.3011 | -13.0848
7.435 | -34.3014 | -13.094
7.4411 | -34.3017 | -13.1041
7.4472 | -34.3019 | -13.1139
7.4533 | -34.3021 | -13.1233
7.4594 | -34.3024 | -13.133
7.4655 | -34.3027 | -13.1429
7.4716 | -34.303 | -13.1526
7.4777 | -34.3033 | -13.1621
7.4838 | -34.3036 | -13.1721
7.4899 | -34.304 | -13.1814
7.496 | -34.3044 | -13.1909
7.5021 | -34.3047 | -13.2009
7.5082 | -34.305 | -13.2105
7.5143 | -34.3053 | -13.2206
7.5204 | -34.3056 | -13.2306
7.5266 | -34.3058 | -13.2402
7.5327 | -34.3061 | -13.2495
7.5388 | -34.3064 | -13.2594
7.5449 | -34.3067 | -13.2689
7.551 | -34.307 | -13.2786
7.5571 | -34.3073 | -13.2881
7.5632 | -34.3077 | -13.2977
7.5693 | -34.308 | -13.3072
7.5754 | -34.3084 | -13.3164
7.5815 | -34.3086 | -13.3261
7.5876 | -34.309 | -13.3362
7.5937 | -34.3093 | -13.3462
7.5998 | -34.3097 | -13.3556
7.6059 | -34.31 | -13.3651
7.612 | -34.3104 | -13.3752
7.6181 | -34.3107 | -13.3846
7.6242 | -34.311 | -13.3944
7.6303 | -34.3113 | -13.4041
7.6364 | -34.3117 | -13.4136
7.6425 | -34.312 | -13.4232
7.6486 | -34.3122 | -13.4325
7.6547 | -34.3126 | -13.4422
7.6608 | -34.313 | -13.4519
7.667 | -34.3132 | -13.4615
7.6731 | -34.3135 | -13.4714
7.6792 | -34.3139 | -13.4812
7.6853 | -34.3142 | -13.491
7.6914 | -34.3145 | -13.5006
7.6975 | -34.3148 | -13.5105
7.7036 | -34.3152 | -13.5201
7.7097 | -34.3155 | -13.5301
7.7158 | -34.3159 | -13.5399
7.7219 | -34.3162 | -13.5495
7.728 | -34.3165 | -13.5594
7.7341 | -34.3169 | -13.569
7.7402 | -34.3172 | -13.5785
7.7463 | -34.3176 | -13.5884
7.7524 | -34.3179 | -13.5984
7.7585 | -34.3182 | -13.608
7.7646 | -34.3185 | -13.6178
7.7707 | -34.3189 | -13.6275
7.7768 | -34.3192 | -13.6373
7.7829 | -34.3196 | -13.6471
7.789 | -34.3199 | -13.6568
7.7951 | -34.3202 | -13.6667
7.8012 | -34.3206 | -13.6763
7.8073 | -34.3211 | -13.6864
7.8135 | -34.3214 | -13.6959
7.8196 | -34.3218 | -13.7056
7.8257 | -34.3221 | -13.7155
7.8318 | -34.3225 | -13.7254
7.8379 | -34.3228 | -13.7352
7.844 | -34.3232 | -13.7451
7.8501 | -34.3236 | -13.7548
7.8562 | -34.3239 | -13.7647
7.8623 | -34.3243 | -13.7745
7.8684 | -34.3246 | -13.7841
7.8745 | -34.325 | -13.7937
7.8806 | -34.3253 | -13.8038
7.8867 | -34.3257 | -13.8131
7.8928 | -34.326 | -13.8225
7.8989 | -34.3264 | -13.8324
7.905 | -34.3267 | -13.8422
7.9111 | -34.3271 | -13.8523
7.9172 | -34.3275 | -13.862
7.9233 | -34.3278 | -13.8718
7.9294 | -34.3282 | -13.8814
7.9355 | -34.3286 | -13.891
7.9416 | -34.3289 | -13.9006
7.9477 | -34.3293 | -13.9102
7.9539 | -34.3297 | -13.9199
7.96 | -34.33 | -13.9298
7.9661 | -34.3304 | -13.9395
7.9722 | -34.3308 | -13.9489
7.9783 | -34.3312 | -13.9591
7.9844 | -34.3316 | -13.9686
7.9905 | -34.3319 | -13.9787
7.9966 | -34.3323 | -13.9885
8.0027 | -34.3327 | -13.9984
8.0088 | -34.333 | -14.008
8.0149 | -34.3334 | -14.0177
8.021 | -34.3338 | -14.0274
8.0271 | -34.3342 | -14.0374
8.0332 | -34.3346 | -14.0469
8.0393 | -34.335 | -14.057
8.0454 | -34.3353 | -14.0672
8.0515 | -34.3359 | -14.0768
8.0576 | -34.3363 | -14.0869
8.0637 | -34.3366 | -14.0968
8.0698 | -34.3369 | -14.1065
8.0759 | -34.3373 | -14.1164
8.082 | -34.3377 | -14.1265
8.0881 | -34.3382 | -14.1364
8.0942 | -34.3386 | -14.1465
8.1004 | -34.339 | -14.1567
8.1065 | -34.3394 | -14.1662
8.1126 | -34.3397 | -14.176
8.1187 | -34.34 | -14.1858
8.1248 | -34.3404 | -14.1949
8.1309 | -34.3407 | -14.2043
8.137 | -34.341 | -14.214
8.1431 | -34.3415 | -14.2235
8.1492 | -34.3419 | -14.2335
8.1553 | -34.3423 | -14.2432
8.1614 | -34.3427 | -14.2528
8.1675 | -34.3431 | -14.2622
8.1736 | -34.3435 | -14.2717
8.1797 | -34.3439 | -14.2819
8.1858 | -34.3443 | -14.292
8.1919 | -34.3448 | -14.3017
8.198 | -34.3452 | -14.3113
8.2041 | -34.3457 | -14.3212
8.2102 | -34.3461 | -14.331
8.2163 | -34.3464 | -14.3409
8.2224 | -34.3469 | -14.351
8.2285 | -34.3473 | -14.3612
8.2346 | -34.3478 | -14.3707
8.2408 | -34.3482 | -14.3804
8.2469 | -34.3486 | -14.3904
8.253 | -34.3491 | -14.4002
8.2591 | -34.3495 | -14.4101
8.2652 | -34.3501 | -14.4203
8.2713 | -34.3505 | -14.4299
8.2774 | -34.3509 | -14.4399
8.2835 | -34.3513 | -14.4499
8.2896 | -34.3518 | -14.4597
8.2957 | -34.3522 | -14.4695
8.3018 | -34.3527 | -14.4793
8.3079 | -34.3531 | -14.4893
8.314 | -34.3536 | -14.4992
8.3201 | -34.3541 | -14.5089
8.3262 | -34.3545 | -14.5191
8.3323 | -34.3549 | -14.5291
8.3384 | -34.3554 | -14.5389
8.3445 | -34.3558 | -14.5489
8.3506 | -34.3562 | -14.5586
8.3567 | -34.3566 | -14.5684
8.3628 | -34.3571 | -14.5786
8.3689 | -34.3575 | -14.5885
8.375 | -34.358 | -14.5984
8.3812 | -34.3584 | -14.6081
8.3873 | -34.3589 | -14.6182
8.3934 | -34.3593 | -14.6283
8.3995 | -34.3597 | -14.6378
8.4056 | -34.3602 | -14.6479
8.4117 | -34.3606 | -14.6575
8.4178 | -34.3611 | -14.6677
8.4239 | -34.3615 | -14.6776
8.43 | -34.3619 | -14.6876
8.4361 | -34.3624 | -14.6978
8.4422 | -34.3628 | -14.7076
8.4483 | -34.3632 | -14.7176
8.4544 | -34.3637 | -14.7276
8.4605 | -34.3642 | -14.7374
8.4666 | -34.3646 | -14.7477
8.4727 | -34.3651 | -14.7577
8.4788 | -34.3655 | -14.7677
8.4849 | -34.3659 | -14.7775
8.491 | -34.3664 | -14.7874
8.4971 | -34.3669 | -14.7971
8.5032 | -34.3673 | -14.8071
8.5093 | -34.3678 | -14.8169
8.5154 | -34.3682 | -14.827
8.5215 | -34.3687 | -14.8367
8.5277 | -34.3691 | -14.8466
8.5338 | -34.3696 | -14.8565
8.5399 | -34.37 | -14.8662
8.546 | -34.3706 | -14.8759
8.5521 | -34.371 | -14.886
8.5582 | -34.3714 | -14.8958
8.5643 | -34.3719 | -14.9056
8.5704 | -34.3723 | -14.9153
8.5765 | -34.3728 | -14.9254
8.5826 | -34.3733 | -14.9355
8.5887 | -34.3738 | -14.9456
8.5948 | -34.3743 | -14.9559
8.6009 | -34.3747 | -14.9653
8.607 | -34.3752 | -14.9752
8.6131 | -34.3756 | -14.985
8.6192 | -34.3761 | -14.995
8.6253 | -34.3766 | -15.0048
8.6314 | -34.377 | -15.0149
8.6375 | -34.3775 | -15.0247
8.6436 | -34.3779 | -15.0345
8.6497 | -34.3784 | -15.0442
8.6558 | -34.3788 | -15.0541
8.6619 | -34.3793 | -15.0642
8.6681 | -34.3797 | -15.0743
8.6742 | -34.3802 | -15.0842
8.6803 | -34.3807 | -15.0939
8.6864 | -34.3812 | -15.1035
8.6925 | -34.3816 | -15.1134
8.6986 | -34.382 | -15.1232
8.7047 | -34.3825 | -15.1332
8.7108 | -34.3829 | -15.1432
8.7169 | -34.3834 | -15.153
8.723 | -34.3838 | -15.163
8.7291 | -34.3843 | -15.1727
8.7352 | -34.3847 | -15.1829
8.7413 | -34.3852 | -15.1931
8.7474 | -34.3856 | -15.2027
8.7535 | -34.3861 | -15.2126
8.7596 | -34.3866 | -15.2226
8.7657 | -34.387 | -15.2328
8.7718 | -34.3875 | -15.2429
8.7779 | -34.388 | -15.2525
8.784 | -34.3884 | -15.2625
8.7901 | -34.389 | -15.2727
8.7962 | -34.3894 | -15.2824
8.8023 | -34.3898 | -15.2922
8.8084 | -34.3903 | -15.3023
8.8146 | -34.3908 | -15.3124
8.8207 | -34.3912 | -15.3223
8.8268 | -34.3917 | -15.3319
8.8329 | -34.3921 | -15.3419
8.839 | -34.3927 | -15.3519
8.8451 | -34.3931 | -15.3618
8.8512 | -34.3935 | -15.3713
8.8573 | -34.394 | -15.3814
8.8634 | -34.3945 | -15.3917
8.8695 | -34.3949 | -15.4018
8.8756 | -34.3954 | -15.412
8.8817 | -34.3958 | -15.4218
8.8878 | -34.3963 | -15.4319
8.8939 | -34.3967 | -15.4419
8.9 | -34.3972 | -15.4519
8.9061 | -34.3976 | -15.4617
8.9122 | -34.3981 | -15.4717
8.9183 | -34.3986 | -15.4817
8.9244 | -34.3991 | -15.492
8.9305 | -34.3995 | -15.5016
8.9366 | -34.3999 | -15.5114
8.9427 | -34.4004 | -15.5212
8.9488 | -34.4008 | -15.5312
8.955 | -34.4012 | -15.5409
8.9611 | -34.4017 | -15.5512
8.9672 | -34.4021 | -15.5611
8.9733 | -34.4026 | -15.5708
8.9794 | -34.4031 | -15.5804
8.9855 | -34.4036 | -15.5905
8.9916 | -34.4041 | -15.6007
8.9977 | -34.4045 | -15.6103
9.0038 | -34.405 | -15.6204
9.0099 | -34.4055 | -15.6305
9.016 | -34.4059 | -15.6403
9.0221 | -34.4064 | -15.6503
9.0282 | -34.4069 | -15.6604
9.0343 | -34.4074 | -15.6702
9.0404 | -34.4078 | -15.6802
9.0465 | -34.4083 | -15.6904
9.0526 | -34.4088 | -15.7004
9.0587 | -34.4093 | -15.7107
9.0648 | -34.4098 | -15.7204
9.0709 | -34.4102 | -15.7306
9.077 | -34.4107 | -15.7403
9.0831 | -34.4111 | -15.7505
9.0892 | -34.4115 | -15.7604
9.0953 | -34.412 | -15.7704
9.1015 | -34.4124 | -15.7807
9.1076 | -34.4129 | -15.7907
9.1137 | -34.4134 | -15.8005
9.1198 | -34.4138 | -15.8103
9.1259 | -34.4142 | -15.8198
9.132 | -34.4147 | -15.8296
9.1381 | -34.4151 | -15.8395
9.1442 | -34.4155 | -15.8491
9.1503 | -34.4159 | -15.8594
9.1564 | -34.4163 | -15.8693
9.1625 | -34.4167 | -15.8789
9.1686 | -34.4173 | -15.8889
9.1747 | -34.4177 | -15.8987
9.1808 | -34.4182 | -15.9089
9.1869 | -34.4187 | -15.9188
9.193 | -34.4191 | -15.9287
9.1991 | -34.4195 | -15.9388
9.2052 | -34.42 | -15.9485
9.2113 | -34.4205 | -15.9582
9.2174 | -34.421 | -15.9681
9.2235 | -34.4214 | -15.9777
9.2296 | -34.4219 | -15.9877
9.2357 | -34.4223 | -15.9975
9.2419 | -34.4228 | -16.0071
9.248 | -34.4232 | -16.0173
9.2541 | -34.4237 | -16.0272
9.2602 | -34.4241 | -16.0374
9.2663 | -34.4246 | -16.047
9.2724 | -34.425 | -16.057
9.2785 | -34.4255 | -16.0669
9.2846 | -34.4259 | -16.0768
9.2907 | -34.4264 | -16.0867
9.2968 | -34.4268 | -16.0968
9.3029 | -34.4272 | -16.1066
9.309 | -34.4277 | -16.1164
9.3151 | -34.4281 | -16.126
9.3212 | -34.4284 | -16.1358
9.3273 | -34.4289 | -16.1453
9.3334 | -34.4293 | -16.1549
9.3395 | -34.4297 | -16.1649
9.3456 | -34.4301 | -16.1749
9.3517 | -34.4305 | -16.1848
9.3578 | -34.4309 | -16.1943
9.3639 | -34.4314 | -16.2039
9.37 | -34.4319 | -16.2139
9.3761 | -34.4323 | -16.2235
9.3822 | -34.4327 | -16.2337
9.3884 | -34.4332 | -16.2425
9.3945 | -34.4336 | -16.2523
9.4006 | -34.4341 | -16.2617
9.4067 | -34.4346 | -16.2714
9.4128 | -34.435 | -16.281
9.4189 | -34.4354 | -16.2909
9.425 | -34.4358 | -16.3004
9.4311 | -34.4362 | -16.3098
9.4372 | -34.4367 | -16.3192
9.4433 | -34.4371 | -16.329
9.4494 | -34.4376 | -16.3388
9.4555 | -34.4381 | -16.3488
9.4616 | -34.4384 | -16.3582
9.4677 | -34.4389 | -16.3675
9.4738 | -34.4393 | -16.3773
9.4799 | -34.4397 | -16.3868
9.486 | -34.4401 | -16.3965
9.4921 | -34.4406 | -16.4061
9.4982 | -34.4409 | -16.4146
9.5043 | -34.4413 | -16.4242
9.5104 | -34.4417 | -16.4336
9.5165 | -34.4421 | -16.4431
9.5226 | -34.4425 | -16.4528
9.5288 | -34.4429 | -16.4615
9.5349 | -34.4433 | -16.471
9.541 | -34.4438 | -16.4809
9.5471 | -34.4442 | -16.4901
9.5532 | -34.4446 | -16.4998
9.5593 | -34.4451 | -16.5098
9.5654 | -34.4455 | -16.5193
9.5715 | -34.4458 | -16.5285
9.5776 | -34.4462 | -16.5383
9.5837 | -34.4466 | -16.5479
9.5898 | -34.4469 | -16.5577
9.5959 | -34.4474 | -16.5673
9.602 | -34.4478 | -16.5762
9.6081 | -34.4483 | -16.5853
9.6142 | -34.4487 | -16.5946
9.6203 | -34.4491 | -16.6043
9.6264 | -34.4495 | -16.614
9.6325 | -34.4499 | -16.6235
9.6386 | -34.4503 | -16.6326
9.6447 | -34.4507 | -16.642
9.6508 | -34.4511 | -16.6518
9.6569 | -34.4514 | -16.6615
9.663 | -34.4519 | -16.6712
9.6691 | -34.4523 | -16.6798
9.6753 | -34.4527 | -16.689
9.6814 | -34.4531 | -16.6981
9.6875 | -34.4536 | -16.7073
9.6936 | -34.454 | -16.7165
9.6997 | -34.4544 | -16.7262
9.7058 | -34.4548 | -16.7349
9.7119 | -34.4553 | -16.7442
9.718 | -34.4557 | -16.7538
9.7241 | -34.4561 | -16.7633
9.7302 | -34.4566 | -16.7729
9.7363 | -34.4569 | -16.7818
9.7424 | -34.4573 | -16.7908
9.7485 | -34.4577 | -16.8003
9.7546 | -34.4581 | -16.809
9.7607 | -34.4585 | -16.8186
9.7668 | -34.4589 | -16.8278
9.7729 | -34.4593 | -16.837
9.779 | -34.4597 | -16.8456
9.7851 | -34.4601 | -16.8544
9.7912 | -34.4605 | -16.8634
9.7973 | -34.4609 | -16.8728
9.8034 | -34.4613 | -16.8821
9.8095 | -34.4617 | -16.8916
9.8157 | -34.4621 | -16.9006
9.8218 | -34.4624 | -16.9101
9.8279 | -34.4628 | -16.9192
9.834 | -34.4631 | -16.9288
9.8401 | -34.4636 | -16.9381
9.8462 | -34.464 | -16.9467
9.8523 | -34.4643 | -16.956
9.8584 | -34.4647 | -16.9656
9.8645 | -34.4651 | -16.9751
9.8706 | -34.4655 | -16.9844
9.8767 | -34.4659 | -16.9936
9.8828 | -34.4663 | -17.0024
9.8889 | -34.4667 | -17.0115
9.895 | -34.4671 | -17.0207
9.9011 | -34.4676 | -17.0299
9.9072 | -34.468 | -17.0392
9.9133 | -34.4685 | -17.0481
9.9194 | -34.4688 | -17.0569
9.9255 | -34.4692 | -17.0659
9.9316 | -34.4696 | -17.0751
9.9377 | -34.4699 | -17.0842
9.9438 | -34.4703 | -17.0932
9.9499 | -34.4707 | -17.1024
9.956 | -34.471 | -17.1113
9.9622 | -34.4714 | -17.1205
9.9683 | -34.4718 | -17.1295
9.9744 | -34.4722 | -17.1384
9.9805 | -34.4726 | -17.1476
9.9866 | -34.473 | -17.1563
9.9927 | -34.4733 | -17.1653
9.9988 | -34.4737 | -17.174
10.0049 | -34.4741 | -17.1832
10.011 | -34.4745 | -17.1922
10.0171 | -34.4749 | -17.2011
10.0232 | -34.4753 | -17.2096
10.0293 | -34.4757 | -17.2188
10.0354 | -34.4761 | -17.2278
10.0415 | -34.4765 | -17.2367
10.0476 | -34.4769 | -17.2461
10.0537 | -34.4773 | -17.2552
10.0598 | -34.4777 | -17.2637
10.0659 | -34.4781 | -17.2728
10.072 | -34.4784 | -17.2815
10.0781 | -34.4788 | -17.2905
10.0842 | -34.4792 | -17.2996
10.0903 | -34.4796 | -17.3086
10.0964 | -34.48 | -17.317
10.1026 | -34.4803 | -17.3261
10.1087 | -34.4807 | -17.3351
10.1148 | -34.4811 | -17.344
10.1209 | -34.4815 | -17.3529
10.127 | -34.4819 | -17.3614
10.1331 | -34.4823 | -17.3704
10.1392 | -34.4826 | -17.379
10.1453 | -34.483 | -17.3879
10.1514 | -34.4834 | -17.3967
10.1575 | -34.4838 | -17.4058
10.1636 | -34.4841 | -17.4141
10.1697 | -34.4846 | -17.4231
10.1758 | -34.4849 | -17.4317
10.1819 | -34.4853 | -17.4404
10.188 | -34.4857 | -17.4495
10.1941 | -34.4861 | -17.4581
10.2002 | -34.4865 | -17.4665
10.2063 | -34.4868 | -17.4753
10.2124 | -34.4873 | -17.4841
10.2185 | -34.4876 | -17.4931
10.2246 | -34.488 | -17.5018
10.2307 | -34.4884 | -17.5107
10.2368 | -34.4887 | -17.5186
10.2429 | -34.4891 | -17.5274
10.2491 | -34.4895 | -17.5364
10.2552 | -34.4898 | -17.5449
10.2613 | -34.4902 | -17.5538
10.2674 | -34.4906 | -17.5624
10.2735 | -34.491 | -17.5714
10.2796 | -34.4915 | -17.5799
10.2857 | -34.4918 | -17.589
10.2918 | -34.4922 | -17.5979
10.2979 | -34.4927 | -17.6066
10.304 | -34.4931 | -17.6152
10.3101 | -34.4934 | -17.624
10.3162 | -34.4939 | -17.6328
10.3223 | -34.4942 | -17.6413
10.3284 | -34.4946 | -17.6501
10.3345 | -34.495 | -17.659
10.3406 | -34.4954 | -17.6675
10.3467 | -34.4957 | -17.6761
10.3528 | -34.4962 | -17.6848
10.3589 | -34.4965 | -17.6933
10.365 | -34.4969 | -17.7023
10.3711 | -34.4973 | -17.7112
10.3772 | -34.4977 | -17.7201
10.3833 | -34.4981 | -17.7286
10.3895 | -34.4985 | -17.7371
10.3956 | -34.4988 | -17.746
10.4017 | -34.4992 | -17.7546
10.4078 | -34.4996 | -17.7634
10.4139 | -34.5 | -17.772
10.42 | -34.5004 | -17.7806
10.4261 | -34.5008 | -17.7894
10.4322 | -34.5012 | -17.7979
10.4383 | -34.5016 | -17.8064
10.4444 | -34.502 | -17.8149
10.4505 | -34.5023 | -17.8237
10.4566 | -34.5027 | -17.8321
10.4627 | -34.5031 | -17.8407
10.4688 | -34.5035 | -17.849
10.4749 | -34.5039 | -17.858
10.481 | -34.5043 | -17.8668
10.4871 | -34.5047 | -17.8752
10.4932 | -34.505 | -17.8837
10.4993 | -34.5054 | -17.8926
10.5054 | -34.5058 | -17.9011
10.5115 | -34.5062 | -17.9099
10.5176 | -34.5066 | -17.9184
10.5237 | -34.507 | -17.9271
10.5298 | -34.5074 | -17.9355
10.536 | -34.5078 | -17.9442
10.5421 | -34.5082 | -17.9529
10.5482 | -34.5086 | -17.9615
10.5543 | -34.509 | -17.9702
10.5604 | -34.5094 | -17.9785
10.5665 | -34.5098 | -17.9871
10.5726 | -34.5102 | -17.9959
10.5787 | -34.5106 | -18.0044
10.5848 | -34.511 | -18.0128
10.5909 | -34.5114 | -18.0216
10.597 | -34.5119 | -18.0306
10.6031 | -34.5123 | -18.0393
10.6092 | -34.5127 | -18.0481
10.6153 | -34.513 | -18.0567
10.6214 | -34.5135 | -18.0652
10.6275 | -34.5138 | -18.074
10.6336 | -34.5142 | -18.0829
10.6397 | -34.5146 | -18.0914
10.6458 | -34.5149 | -18.1
10.6519 | -34.5153 | -18.1089
10.658 | -34.5157 | -18.1178
10.6641 | -34.516 | -18.1266
10.6702 | -34.5164 | -18.1351
10.6764 | -34.5168 | -18.1439
10.6825 | -34.5172 | -18.1521
10.6886 | -34.5176 | -18.1606
10.6947 | -34.518 | -18.1694
10.7008 | -34.5184 | -18.1782
10.7069 | -34.5189 | -18.1869
10.713 | -34.5193 | -18.196
10.7191 | -34.5197 | -18.2046
10.7252 | -34.5201 | -18.2133
10.7313 | -34.5205 | -18.2223
10.7374 | -34.5209 | -18.2311
10.7435 | -34.5212 | -18.2399
10.7496 | -34.5217 | -18.2487
10.7557 | -34.522 | -18.2574
10.7618 | -34.5224 | -18.2662
10.7679 | -34.5228 | -18.2748
10.774 | -34.5232 | -18.2839
10.7801 | -34.5236 | -18.2928
10.7862 | -34.524 | -18.3012
10.7923 | -34.5243 | -18.31
10.7984 | -34.5248 | -18.319
10.8045 | -34.5252 | -18.3275
10.8106 | -34.5256 | -18.3366
10.8168 | -34.5259 | -18.3454
10.8229 | -34.5263 | -18.3539
10.829 | -34.5267 | -18.3628
10.8351 | -34.5272 | -18.3716
10.8412 | -34.5276 | -18.3804
10.8473 | -34.528 | -18.3892
10.8534 | -34.5284 | -18.398
10.8595 | -34.5288 | -18.407
10.8656 | -34.5292 | -18.4158
10.8717 | -34.5296 | -18.4248
10.8778 | -34.5301 | -18.4339
10.8839 | -34.5304 | -18.4426
10.89 | -34.5308 | -18.4513
10.8961 | -34.5312 | -18.4603
10.9022 | -34.5317 | -18.4692
10.9083 | -34.5322 | -18.4781
10.9144 | -34.5325 | -18.4866
10.9205 | -34.5329 | -18.4956
10.9266 | -34.5333 | -18.5044
10.9327 | -34.5338 | -18.5132
10.9388 | -34.5342 | -18.5222
10.9449 | -34.5347 | -18.5309
10.951 | -34.5351 | -18.5396
10.9571 | -34.5355 | -18.5486
10.9633 | -34.5359 | -18.5573
10.9694 | -34.5363 | -18.5662
10.9755 | -34.5368 | -18.5752
10.9816 | -34.5372 | -18.5842
10.9877 | -34.5376 | -18.5928
10.9938 | -34.538 | -18.602
10.9999 | -34.5385 | -18.6107
11.006 | -34.5389 | -18.6196
11.0121 | -34.5393 | -18.6285
11.0182 | -34.5397 | -18.6374
11.0243 | -34.5401 | -18.6463
11.0304 | -34.5405 | -18.6552
11.0365 | -34.5409 | -18.6639
11.0426 | -34.5414 | -18.6731
11.0487 | -34.5418 | -18.682
11.0548 | -34.5422 | -18.6909
11.0609 | -34.5426 | -18.6999
11.067 | -34.5431 | -18.7089
11.0731 | -34.5435 | -18.718
11.0792 | -34.544 | -18.7271
11.0853 | -34.5444 | -18.7359
11.0914 | -34.5448 | -18.7451
11.0975 | -34.5452 | -18.7539
11.1037 | -34.5456 | -18.7627
11.1098 | -34.546 | -18.772
11.1159 | -34.5465 | -18.7813
11.122 | -34.5469 | -18.7904
11.1281 | -34.5473 | -18.7992
11.1342 | -34.5477 | -18.8081
11.1403 | -34.5482 | -18.817
11.1464 | -34.5486 | -18.8261
11.1525 | -34.549 | -18.8352
11.1586 | -34.5494 | -18.8442
11.1647 | -34.5498 | -18.8531
11.1708 | -34.5502 | -18.8622
11.1769 | -34.5507 | -18.8712
11.183 | -34.5511 | -18.8805
11.1891 | -34.5516 | -18.8896
11.1952 | -34.552 | -18.8987
11.2013 | -34.5524 | -18.9078
11.2074 | -34.5529 | -18.9167
11.2135 | -34.5533 | -18.926
11.2196 | -34.5537 | -18.9352
11.2257 | -34.5542 | -18.9444
11.2318 | -34.5546 | -18.9534
11.2379 | -34.5551 | -18.9625
11.244 | -34.5555 | -18.9717
11.2502 | -34.556 | -18.9806
11.2563 | -34.5565 | -18.99
11.2624 | -34.5569 | -18.9991
11.2685 | -34.5573 | -19.0082
11.2746 | -34.5577 | -19.0175
11.2807 | -34.5581 | -19.0267
11.2868 | -34.5586 | -19.0359
11.2929 | -34.559 | -19.0451
11.299 | -34.5595 | -19.0545
11.3051 | -34.5599 | -19.0636
11.3112 | -34.5603 | -19.0729
11.3173 | -34.5608 | -19.0817
11.3234 | -34.5613 | -19.091
11.3295 | -34.5617 | -19.1003
11.3356 | -34.5622 | -19.1097
11.3417 | -34.5626 | -19.1186
11.3478 | -34.563 | -19.1279
11.3539 | -34.5635 | -19.1373
11.36 | -34.5639 | -19.1467
11.3661 | -34.5644 | -19.1557
11.3722 | -34.5648 | -19.1648
11.3783 | -34.5653 | -19.174
11.3844 | -34.5657 | -19.1833
11.3906 | -34.5661 | -19.1926
11.3967 | -34.5666 | -19.2019
11.4028 | -34.567 | -19.2113
11.4089 | -34.5675 | -19.2204
11.415 | -34.5679 | -19.2296
11.4211 | -34.5683 | -19.2391
11.4272 | -34.5687 | -19.2481
11.4333 | -34.5692 | -19.2576
11.4394 | -34.5696 | -19.2666
11.4455 | -34.57 | -19.2756
11.4516 | -34.5705 | -19.2851
11.4577 | -34.5709 | -19.2942
11.4638 | -34.5714 | -19.3037
11.4699 | -34.5718 | -19.3129
11.476 | -34.5722 | -19.3224
11.4821 | -34.5727 | -19.3316
11.4882 | -34.5731 | -19.3405
11.4943 | -34.5735 | -19.3497
11.5004 | -34.574 | -19.359
11.5065 | -34.5745 | -19.3685
11.5126 | -34.575 | -19.3775
11.5187 | -34.5754 | -19.3865
11.5248 | -34.5758 | -19.3959
11.5309 | -34.5763 | -19.4053
11.5371 | -34.5768 | -19.4148
11.5432 | -34.5772 | -19.4242
11.5493 | -34.5777 | -19.4339
11.5554 | -34.5782 | -19.4435
11.5615 | -34.5787 | -19.4531
11.5676 | -34.5792 | -19.4623
11.5737 | -34.5796 | -19.4719
11.5798 | -34.58 | -19.4814
11.5859 | -34.5805 | -19.4906
11.592 | -34.5809 | -19.4999
11.5981 | -34.5813 | -19.5093
11.6042 | -34.5818 | -19.5188
11.6103 | -34.5822 | -19.528
11.6164 | -34.5827 | -19.5374
11.6225 | -34.5831 | -19.5468
11.6286 | -34.5836 | -19.5559
11.6347 | -34.584 | -19.5652
11.6408 | -34.5844 | -19.5746
11.6469 | -34.5849 | -19.584
11.653 | -34.5853 | -19.5934
11.6591 | -34.5858 | -19.6028
11.6652 | -34.5862 | -19.612
11.6713 | -34.5866 | -19.6212
11.6775 | -34.5871 | -19.6306
11.6836 | -34.5876 | -19.6397
11.6897 | -34.588 | -19.649
11.6958 | -34.5885 | -19.6582
11.7019 | -34.5889 | -19.6676
11.708 | -34.5893 | -19.677
11.7141 | -34.5898 | -19.6864
11.7202 | -34.5902 | -19.6956
11.7263 | -34.5906 | -19.7049
11.7324 | -34.5911 | -19.7144
11.7385 | -34.5915 | -19.7237
11.7446 | -34.5919 | -19.733
11.7507 | -34.5924 | -19.7424
11.7568 | -34.5928 | -19.7516
11.7629 | -34.5932 | -19.7611
11.769 | -34.5937 | -19.7702
11.7751 | -34.5941 | -19.7796
11.7812 | -34.5946 | -19.789
11.7873 | -34.595 | -19.7984
11.7934 | -34.5955 | -19.808
11.7995 | -34.596 | -19.8171
11.8056 | -34.5964 | -19.8265
11.8117 | -34.5969 | -19.8356
11.8178 | -34.5973 | -19.845
11.824 | -34.5978 | -19.8543
11.8301 | -34.5982 | -19.8638
11.8362 | -34.5987 | -19.8728
11.8423 | -34.5991 | -19.8819
11.8484 | -34.5996 | -19.8914
11.8545 | -34.6001 | -19.9009
11.8606 | -34.6005 | -19.9102
11.8667 | -34.601 | -19.9196
11.8728 | -34.6014 | -19.9289
11.8789 | -34.6019 | -19.9383
11.885 | -34.6024 | -19.9474
11.8911 | -34.6028 | -19.9569
11.8972 | -34.6033 | -19.9663
11.9033 | -34.6037 | -19.9756
11.9094 | -34.6042 | -19.9849
11.9155 | -34.6046 | -19.9943
11.9216 | -34.6051 | -20.0038
11.9277 | -34.6056 | -20.013
11.9338 | -34.606 | -20.0223
11.9399 | -34.6065 | -20.032
11.946 | -34.607 | -20.041
11.9521 | -34.6074 | -20.0505
11.9582 | -34.6079 | -20.0598
11.9644 | -34.6083 | -20.0695
11.9705 | -34.6088 | -20.0786
11.9766 | -34.6092 | -20.0879
11.9827 | -34.6097 | -20.0973
11.9888 | -34.6102 | -20.1066
11.9949 | -34.6107 | -20.1157
12.001 | -34.6111 | -20.1252
12.0071 | -34.6116 | -20.1345
12.0132 | -34.6121 | -20.1436
12.0193 | -34.6125 | -20.153
12.0254 | -34.6129 | -20.1623
12.0315 | -34.6134 | -20.1715
12.0376 | -34.6138 | -20.1806
12.0437 | -34.6143 | -20.19
12.0498 | -34.6148 | -20.1994
12.0559 | -34.6153 | -20.2086
12.062 | -34.6157 | -20.2179
12.0681 | -34.6162 | -20.2274
12.0742 | -34.6166 | -20.2367
12.0803 | -34.6171 | -20.2461
12.0864 | -34.6175 | -20.2555
12.0925 | -34.618 | -20.265
12.0986 | -34.6185 | -20.2744
12.1047 | -34.6189 | -20.2837
12.1109 | -34.6194 | -20.2928
12.117 | -34.6198 | -20.3021
12.1231 | -34.6203 | -20.3113
12.1292 | -34.6207 | -20.3207
12.1353 | -34.6212 | -20.3299
12.1414 | -34.6217 | -20.3395
12.1475 | -34.6221 | -20.3489
12.1536 | -34.6225 | -20.3583
12.1597 | -34.623 | -20.3675
12.1658 | -34.6235 | -20.3768
12.1719 | -34.6239 | -20.3858
12.178 | -34.6244 | -20.3951
12.1841 | -34.6248 | -20.4045
12.1902 | -34.6253 | -20.4138
12.1963 | -34.6258 | -20.423
12.2024 | -34.6263 | -20.4321
12.2085 | -34.6267 | -20.4413
12.2146 | -34.6272 | -20.4505
12.2207 | -34.6276 | -20.4597
12.2268 | -34.6281 | -20.4693
12.2329 | -34.6286 | -20.4787
12.239 | -34.629 | -20.4877
12.2451 | -34.6295 | -20.4972
12.2513 | -34.63 | -20.5066
12.2574 | -34.6304 | -20.5159
12.2635 | -34.6309 | -20.5251
12.2696 | -34.6314 | -20.5341
12.2757 | -34.6318 | -20.5432
12.2818 | -34.6323 | -20.5526
12.2879 | -34.6328 | -20.5619
12.294 | -34.6332 | -20.5712
12.3001 | -34.6337 | -20.5805
12.3062 | -34.6342 | -20.5898
12.3123 | -34.6347 | -20.5989
12.3184 | -34.6351 | -20.608
12.3245 | -34.6356 | -20.617
12.3306 | -34.636 | -20.6262
12.3367 | -34.6365 | -20.6355
12.3428 | -34.6369 | -20.6446
12.3489 | -34.6374 | -20.6535
12.355 | -34.6379 | -20.663
12.3611 | -34.6384 | -20.6723
12.3672 | -34.6388 | -20.6817
12.3733 | -34.6393 | -20.6909
12.3794 | -34.6398 | -20.6998
12.3855 | -34.6402 | -20.7088
12.3916 | -34.6407 | -20.7182
12.3978 | -34.6412 | -20.7274
12.4039 | -34.6417 | -20.7367
12.41 | -34.6421 | -20.7459
12.4161 | -34.6426 | -20.7545
12.4222 | -34.6431 | -20.7638
12.4283 | -34.6435 | -20.7729
12.4344 | -34.644 | -20.7822
12.4405 | -34.6445 | -20.7914
12.4466 | -34.645 | -20.8006
12.4527 | -34.6454 | -20.8096
12.4588 | -34.6459 | -20.8187
12.4649 | -34.6463 | -20.8277
12.471 | -34.6468 | -20.8371
12.4771 | -34.6472 | -20.8463
12.4832 | -34.6477 | -20.8554
12.4893 | -34.6481 | -20.8643
12.4954 | -34.6486 | -20.8732
12.5015 | -34.649 | -20.8823
12.5076 | -34.6494 | -20.8915
12.5137 | -34.6499 | -20.9007
12.5198 | -34.6504 | -20.9097
12.5259 | -34.6509 | -20.9186
12.532 | -34.6513 | -20.9275
12.5382 | -34.6518 | -20.9367
12.5443 | -34.6522 | -20.9458
12.5504 | -34.6527 | -20.955
12.5565 | -34.6532 | -20.964
12.5626 | -34.6536 | -20.9732
12.5687 | -34.6541 | -20.9825
12.5748 | -34.6546 | -20.9917
12.5809 | -34.655 | -21.001
12.587 | -34.6555 | -21.0101
12.5931 | -34.6559 | -21.019
12.5992 | -34.6564 | -21.0279
12.6053 | -34.6568 | -21.0368
12.6114 | -34.6573 | -21.0459
12.6175 | -34.6578 | -21.0551
12.6236 | -34.6582 | -21.0642
12.6297 | -34.6587 | -21.0733
12.6358 | -34.6591 | -21.0822
12.6419 | -34.6596 | -21.0914
12.648 | -34.6601 | -21.1006
12.6541 | -34.6605 | -21.1097
12.6602 | -34.661 | -21.1185
12.6663 | -34.6615 | -21.1277
12.6724 | -34.6619 | -21.1367
12.6785 | -34.6623 | -21.1458
12.6847 | -34.6628 | -21.1546
12.6908 | -34.6633 | -21.1639
12.6969 | -34.6638 | -21.1728
12.703 | -34.6642 | -21.1821
12.7091 | -34.6647 | -21.1912
12.7152 | -34.6651 | -21.2001
12.7213 | -34.6656 | -21.2092
12.7274 | -34.666 | -21.2184
12.7335 | -34.6665 | -21.2271
12.7396 | -34.667 | -21.236
12.7457 | -34.6674 | -21.245
12.7518 | -34.6679 | -21.2541
12.7579 | -34.6684 | -21.2631
12.764 | -34.6688 | -21.2723
12.7701 | -34.6692 | -21.2811
12.7762 | -34.6697 | -21.2898
12.7823 | -34.6701 | -21.2988
12.7884 | -34.6706 | -21.3081
12.7945 | -34.671 | -21.3174
12.8006 | -34.6715 | -21.3263
12.8067 | -34.672 | -21.3351
12.8128 | -34.6724 | -21.3439
12.8189 | -34.6729 | -21.3529
12.8251 | -34.6733 | -21.3618
12.8312 | -34.6738 | -21.3709
12.8373 | -34.6742 | -21.3796
12.8434 | -34.6747 | -21.3888
12.8495 | -34.6752 | -21.3975
12.8556 | -34.6756 | -21.4065
12.8617 | -34.676 | -21.4156
12.8678 | -34.6764 | -21.4247
12.8739 | -34.6769 | -21.4333
12.88 | -34.6773 | -21.4421
12.8861 | -34.6777 | -21.4509
12.8922 | -34.6782 | -21.4599
12.8983 | -34.6787 | -21.4689
12.9044 | -34.6791 | -21.478
12.9105 | -34.6796 | -21.4867
12.9166 | -34.6801 | -21.4957
12.9227 | -34.6806 | -21.5045
12.9288 | -34.681 | -21.5133
12.9349 | -34.6815 | -21.5222
12.941 | -34.682 | -21.5311
12.9471 | -34.6824 | -21.5399
12.9532 | -34.6829 | -21.5489
12.9593 | -34.6834 | -21.5581
12.9654 | -34.6838 | -21.5672
12.9716 | -34.6843 | -21.5763
12.9777 | -34.6848 | -21.5849
12.9838 | -34.6853 | -21.5938
12.9899 | -34.6857 | -21.6028
12.996 | -34.6862 | -21.6118
13.0021 | -34.6867 | -21.6208
13.0082 | -34.6872 | -21.6294
13.0143 | -34.6877 | -21.638
13.0204 | -34.6881 | -21.6466
13.0265 | -34.6886 | -21.6556
13.0326 | -34.6891 | -21.6645
13.0387 | -34.6895 | -21.6738
13.0448 | -34.69 | -21.6826
13.0509 | -34.6904 | -21.6912
13.057 | -34.6909 | -21.7001
13.0631 | -34.6913 | -21.7089
13.0692 | -34.6918 | -21.7179
13.0753 | -34.6923 | -21.7268
13.0814 | -34.6927 | -21.7358
13.0875 | -34.6932 | -21.7444
13.0936 | -34.6936 | -21.7536
13.0997 | -34.694 | -21.7622
13.1058 | -34.6945 | -21.7711
13.112 | -34.695 | -21.7799
13.1181 | -34.6955 | -21.7885
13.1242 | -34.6959 | -21.7974
13.1303 | -34.6964 | -21.8062
13.1364 | -34.6968 | -21.8152
13.1425 | -34.6973 | -21.8241
13.1486 | -34.6977 | -21.8327
13.1547 | -34.6982 | -21.8414
13.1608 | -34.6986 | -21.85
13.1669 | -34.6991 | -21.8586
13.173 | -34.6996 | -21.8675
13.1791 | -34.7 | -21.8762
13.1852 | -34.7005 | -21.885
13.1913 | -34.7009 | -21.8937
13.1974 | -34.7014 | -21.9028
13.2035 | -34.7018 | -21.9115
13.2096 | -34.7022 | -21.9202
13.2157 | -34.7027 | -21.929
13.2218 | -34.7032 | -21.938
13.2279 | -34.7037 | -21.9466
13.234 | -34.7041 | -21.9554
13.2401 | -34.7046 | -21.9644
13.2462 | -34.7051 | -21.9732
13.2524 | -34.7056 | -21.982
13.2585 | -34.706 | -21.9905
13.2646 | -34.7065 | -21.9994
13.2707 | -34.7069 | -22.0083
13.2768 | -34.7074 | -22.017
13.2829 | -34.7079 | -22.0257
13.289 | -34.7084 | -22.0344
13.2951 | -34.7088 | -22.043
13.3012 | -34.7093 | -22.0518
13.3073 | -34.7098 | -22.0607
13.3134 | -34.7102 | -22.0696
13.3195 | -34.7106 | -22.0781
13.3256 | -34.7111 | -22.0869
13.3317 | -34.7115 | -22.0955
13.3378 | -34.712 | -22.1041
13.3439 | -34.7125 | -22.113
13.35 | -34.7129 | -22.1223
13.3561 | -34.7133 | -22.131
13.3622 | -34.7137 | -22.1399
13.3683 | -34.7142 | -22.1484
13.3744 | -34.7147 | -22.1572
13.3805 | -34.7152 | -22.1661
13.3866 | -34.7156 | -22.1748
13.3927 | -34.7161 | -22.1838
13.3989 | -34.7165 | -22.1922
13.405 | -34.717 | -22.2006
13.4111 | -34.7175 | -22.2094
13.4172 | -34.7179 | -22.2185
13.4233 | -34.7184 | -22.2273
13.4294 | -34.7188 | -22.2359
13.4355 | -34.7192 | -22.2443
13.4416 | -34.7197 | -22.2528
13.4477 | -34.7202 | -22.2614
13.4538 | -34.7206 | -22.2702
13.4599 | -34.7211 | -22.2792
13.466 | -34.7216 | -22.2879
13.4721 | -34.722 | -22.2962
13.4782 | -34.7225 | -22.3051
13.4843 | -34.723 | -22.314
13.4904 | -34.7234 | -22.3225
13.4965 | -34.7239 | -22.3317
13.5026 | -34.7243 | -22.3406
13.5087 | -34.7248 | -22.3491
13.5148 | -34.7253 | -22.3579
13.5209 | -34.7258 | -22.3669
13.527 | -34.7262 | -22.3756
13.5331 | -34.7267 | -22.3843
13.5393 | -34.7272 | -22.3927
13.5454 | -34.7276 | -22.4017
13.5515 | -34.728 | -22.4103
13.5576 | -34.7285 | -22.4189
13.5637 | -34.729 | -22.4274
13.5698 | -34.7294 | -22.4363
13.5759 | -34.7298 | -22.4445
13.582 | -34.7303 | -22.4532
13.5881 | -34.7307 | -22.4616
13.5942 | -34.7312 | -22.4703
13.6003 | -34.7317 | -22.4787
13.6064 | -34.7322 | -22.4876
13.6125 | -34.7326 | -22.496
13.6186 | -34.7331 | -22.5045
13.6247 | -34.7336 | -22.513
13.6308 | -34.7341 | -22.5216
13.6369 | -34.7345 | -22.5302
13.643 | -34.7349 | -22.5389
13.6491 | -34.7354 | -22.5472
13.6552 | -34.7359 | -22.5558
13.6613 | -34.7363 | -22.5646
13.6674 | -34.7367 | -22.5732
13.6735 | -34.7372 | -22.5818
13.6796 | -34.7376 | -22.5903
13.6858 | -34.7381 | -22.5988
13.6919 | -34.7385 | -22.6074
13.698 | -34.739 | -22.6157
13.7041 | -34.7394 | -22.6246
13.7102 | -34.74 | -22.6332
13.7163 | -34.7404 | -22.6417
13.7224 | -34.7409 | -22.6502
13.7285 | -34.7414 | -22.6585
13.7346 | -34.7418 | -22.6669
13.7407 | -34.7423 | -22.6757
13.7468 | -34.7428 | -22.6843
13.7529 | -34.7432 | -22.6929
13.759 | -34.7437 | -22.7016
13.7651 | -34.7441 | -22.7104
13.7712 | -34.7446 | -22.719
13.7773 | -34.745 | -22.7275
13.7834 | -34.7455 | -22.736
13.7895 | -34.746 | -22.7446
13.7956 | -34.7464 | -22.7532
13.8017 | -34.7468 | -22.7618
13.8078 | -34.7472 | -22.7706
13.8139 | -34.7477 | -22.779
13.82 | -34.7481 | -22.787
13.8262 | -34.7486 | -22.7955
13.8323 | -34.749 | -22.8039
13.8384 | -34.7495 | -22.8126
13.8445 | -34.75 | -22.8211
13.8506 | -34.7504 | -22.8296
13.8567 | -34.7508 | -22.838
13.8628 | -34.7513 | -22.8464
13.8689 | -34.7518 | -22.8555
13.875 | -34.7522 | -22.8637
13.8811 | -34.7526 | -22.8721
13.8872 | -34.7531 | -22.8806
13.8933 | -34.7536 | -22.8888
13.8994 | -34.754 | -22.8974
13.9055 | -34.7545 | -22.9059
13.9116 | -34.7549 | -22.9146
13.9177 | -34.7554 | -22.9232
13.9238 | -34.7559 | -22.9316
13.9299 | -34.7564 | -22.94
13.936 | -34.7568 | -22.9483
13.9421 | -34.7572 | -22.9566
13.9482 | -34.7577 | -22.9649
13.9543 | -34.7581 | -22.9734
13.9604 | -34.7586 | -22.9819
13.9665 | -34.7591 | -22.99
13.9727 | -34.7596 | -22.9985
13.9788 | -34.7601 | -23.0067
13.9849 | -34.7605 | -23.015
13.991 | -34.7609 | -23.0235
13.9971 | -34.7614 | -23.0318
14.0032 | -34.7619 | -23.0403
14.0093 | -34.7624 | -23.0488
14.0154 | -34.7628 | -23.0569
14.0215 | -34.7633 | -23.0652
14.0276 | -34.7638 | -23.0736
14.0337 | -34.7643 | -23.0821
14.0398 | -34.7647 | -23.0903
14.0459 | -34.7652 | -23.0988
14.052 | -34.7656 | -23.1073
14.0581 | -34.766 | -23.1158
14.0642 | -34.7665 | -23.1242
14.0703 | -34.767 | -23.1325
14.0764 | -34.7674 | -23.1408
14.0825 | -34.7678 | -23.149
14.0886 | -34.7683 | -23.1575
14.0947 | -34.7687 | -23.166
14.1008 | -34.7692 | -23.1744
14.1069 | -34.7697 | -23.1825
14.1131 | -34.7701 | -23.1911
14.1192 | -34.7706 | -23.1992
14.1253 | -34.771 | -23.2073
14.1314 | -34.7715 | -23.2158
14.1375 | -34.772 | -23.224
14.1436 | -34.7725 | -23.2324
14.1497 | -34.773 | -23.241
14.1558 | -34.7734 | -23.2489
14.1619 | -34.7739 | -23.2575
14.168 | -34.7743 | -23.2661
14.1741 | -34.7748 | -23.2745
14.1802 | -34.7753 | -23.2828
14.1863 | -34.7758 | -23.2914
14.1924 | -34.7762 | -23.2996
14.1985 | -34.7767 | -23.3079
14.2046 | -34.7772 | -23.3165
14.2107 | -34.7777 | -23.3246
14.2168 | -34.7782 | -23.3332
14.2229 | -34.7786 | -23.3416
14.229 | -34.7791 | -23.3501
14.2351 | -34.7795 | -23.3585
14.2412 | -34.78 | -23.367
14.2473 | -34.7805 | -23.3752
14.2534 | -34.7809 | -23.3835
14.2596 | -34.7813 | -23.3916
14.2657 | -34.7818 | -23.3998
14.2718 | -34.7822 | -23.4082
14.2779 | -34.7827 | -23.4167
14.284 | -34.7833 | -23.425
14.2901 | -34.7837 | -23.4334
14.2962 | -34.7842 | -23.4417
14.3023 | -34.7847 | -23.4502
14.3084 | -34.7852 | -23.4586
14.3145 | -34.7856 | -23.4667
14.3206 | -34.7861 | -23.4751
14.3267 | -34.7865 | -23.4836
14.3328 | -34.787 | -23.492
14.3389 | -34.7875 | -23.5002
14.345 | -34.788 | -23.5085
14.3511 | -34.7884 | -23.5169
14.3572 | -34.7889 | -23.5252
14.3633 | -34.7894 | -23.5336
14.3694 | -34.7898 | -23.5418
14.3755 | -34.7903 | -23.55
14.3816 | -34.7908 | -23.5583
14.3877 | -34.7913 | -23.5665
14.3938 | -34.7917 | -23.5749
14.4 | -34.7922 | -23.5832
14.4061 | -34.7927 | -23.5915
14.4122 | -34.7931 | -23.5998
14.4183 | -34.7936 | -23.6078
14.4244 | -34.7941 | -23.6163
14.4305 | -34.7946 | -23.6245
14.4366 | -34.7951 | -23.6327
14.4427 | -34.7956 | -23.6412
14.4488 | -34.796 | -23.6495
14.4549 | -34.7965 | -23.6577
14.461 | -34.797 | -23.666
14.4671 | -34.7974 | -23.6742
14.4732 | -34.7979 | -23.6827
14.4793 | -34.7983 | -23.6909
14.4854 | -34.7988 | -23.6994
14.4915 | -34.7993 | -23.7072
14.4976 | -34.7998 | -23.7154
14.5037 | -34.8002 | -23.7236
14.5098 | -34.8007 | -23.7317
14.5159 | -34.8011 | -23.7401
14.522 | -34.8016 | -23.7484
14.5281 | -34.8021 | -23.7565
14.5342 | -34.8025 | -23.7649
14.5403 | -34.8031 | -23.7733
14.5465 | -34.8035 | -23.7817
14.5526 | -34.804 | -23.7902
14.5587 | -34.8045 | -23.7986
14.5648 | -34.805 | -23.8068
14.5709 | -34.8054 | -23.8149
14.577 | -34.8058 | -23.8231
14.5831 | -34.8063 | -23.8317
14.5892 | -34.8068 | -23.8397
14.5953 | -34.8073 | -23.8479
14.6014 | -34.8077 | -23.8561
14.6075 | -34.8082 | -23.8646
14.6136 | -34.8087 | -23.8728
14.6197 | -34.8091 | -23.8808
14.6258 | -34.8096 | -23.889
14.6319 | -34.8101 | -23.8969
14.638 | -34.8105 | -23.9051
14.6441 | -34.811 | -23.9134
14.6502 | -34.8115 | -23.9218
14.6563 | -34.812 | -23.93
14.6624 | -34.8125 | -23.9382
14.6685 | -34.813 | -23.9461
14.6746 | -34.8134 | -23.9546
14.6807 | -34.8139 | -23.9629
14.6869 | -34.8144 | -23.9711
14.693 | -34.8148 | -23.9794
14.6991 | -34.8153 | -23.9877
14.7052 | -34.8158 | -23.9958
14.7113 | -34.8162 | -24.0042
14.7174 | -34.8167 | -24.0123
14.7235 | -34.8172 | -24.0204
14.7296 | -34.8176 | -24.0286
14.7357 | -34.8181 | -24.0367
14.7418 | -34.8186 | -24.045
14.7479 | -34.8191 | -24.0529
14.754 | -34.8196 | -24.0611
14.7601 | -34.8201 | -24.0693
14.7662 | -34.8205 | -24.0774
14.7723 | -34.821 | -24.0856
14.7784 | -34.8215 | -24.0937
14.7845 | -34.822 | -24.1019
14.7906 | -34.8224 | -24.1102
14.7967 | -34.8229 | -24.1183
14.8028 | -34.8233 | -24.1263
14.8089 | -34.8238 | -24.1344
14.815 | -34.8243 | -24.1429
14.8211 | -34.8248 | -24.1511
14.8272 | -34.8253 | -24.1595
14.8334 | -34.8258 | -24.1677
14.8395 | -34.8262 | -24.1758
14.8456 | -34.8267 | -24.184
14.8517 | -34.8272 | -24.1923
14.8578 | -34.8277 | -24.2005
14.8639 | -34.8282 | -24.2089
14.87 | -34.8287 | -24.2172
14.8761 | -34.8292 | -24.2254
14.8822 | -34.8296 | -24.2337
14.8883 | -34.8301 | -24.242
14.8944 | -34.8305 | -24.2502
14.9005 | -34.831 | -24.2583
14.9066 | -34.8315 | -24.2665
14.9127 | -34.832 | -24.2745
14.9188 | -34.8325 | -24.2828
14.9249 | -34.8329 | -24.2909
14.931 | -34.8334 | -24.2991
14.9371 | -34.8339 | -24.3073
14.9432 | -34.8343 | -24.3156
14.9493 | -34.8348 | -24.3235
14.9554 | -34.8353 | -24.3317
14.9615 | -34.8357 | -24.3399
14.9676 | -34.8362 | -24.3483
14.9738 | -34.8367 | -24.3562
14.9799 | -34.8372 | -24.3647
14.986 | -34.8377 | -24.373
14.9921 | -34.8382 | -24.3812
14.9982 | -34.8386 | -24.3896
15.0043 | -34.8391 | -24.3978
15.0104 | -34.8396 | -24.4061
15.0165 | -34.84 | -24.4144
15.0226 | -34.8406 | -24.4225
15.0287 | -34.841 | -24.4307
15.0348 | -34.8415 | -24.439
15.0409 | -34.842 | -24.4474
15.047 | -34.8425 | -24.4558
15.0531 | -34.843 | -24.4636
15.0592 | -34.8435 | -24.4717
15.0653 | -34.844 | -24.48
15.0714 | -34.8445 | -24.4882
15.0775 | -34.845 | -24.4965
15.0836 | -34.8455 | -24.5046
15.0897 | -34.846 | -24.5128
15.0958 | -34.8465 | -24.5212
15.1019 | -34.847 | -24.5294
15.108 | -34.8475 | -24.5375
15.1141 | -34.8479 | -24.5456
15.1203 | -34.8484 | -24.554
15.1264 | -34.8489 | -24.5624
15.1325 | -34.8494 | -24.5706
15.1386 | -34.8499 | -24.5788
15.1447 | -34.8504 | -24.5868
15.1508 | -34.8509 | -24.5948
15.1569 | -34.8514 | -24.603
15.163 | -34.8519 | -24.6111
15.1691 | -34.8524 | -24.6194
15.1752 | -34.8529 | -24.6276
15.1813 | -34.8534 | -24.6357
15.1874 | -34.8538 | -24.6438
15.1935 | -34.8543 | -24.6517
15.1996 | -34.8548 | -24.66
15.2057 | -34.8553 | -24.668
15.2118 | -34.8558 | -24.6761
15.2179 | -34.8563 | -24.6842
15.224 | -34.8567 | -24.6924
15.2301 | -34.8572 | -24.7006
15.2362 | -34.8577 | -24.7087
15.2423 | -34.8582 | -24.7168
15.2484 | -34.8587 | -24.725
15.2545 | -34.8592 | -24.7333
15.2607 | -34.8597 | -24.7413
15.2668 | -34.8603 | -24.7491
15.2729 | -34.8608 | -24.7571
15.279 | -34.8613 | -24.7653
15.2851 | -34.8618 | -24.7733
15.2912 | -34.8623 | -24.7814
15.2973 | -34.8628 | -24.7894
15.3034 | -34.8633 | -24.7975
15.3095 | -34.8638 | -24.8057
15.3156 | -34.8643 | -24.8137
15.3217 | -34.8647 | -24.8218
15.3278 | -34.8652 | -24.8301
15.3339 | -34.8658 | -24.8382
15.34 | -34.8663 | -24.8462
15.3461 | -34.8668 | -24.8541
15.3522 | -34.8672 | -24.8621
15.3583 | -34.8677 | -24.8704
15.3644 | -34.8682 | -24.8784
15.3705 | -34.8687 | -24.8861
15.3766 | -34.8692 | -24.8943
15.3827 | -34.8697 | -24.9026
15.3888 | -34.8701 | -24.9107
15.3949 | -34.8706 | -24.9189
15.401 | -34.8711 | -24.927
15.4072 | -34.8716 | -24.9348
15.4133 | -34.8721 | -24.9428
15.4194 | -34.8726 | -24.951
15.4255 | -34.8731 | -24.9592
15.4316 | -34.8736 | -24.9673
15.4377 | -34.8741 | -24.9756
15.4438 | -34.8746 | -24.9836
15.4499 | -34.8751 | -24.9917
15.456 | -34.8756 | -24.9996
15.4621 | -34.8761 | -25.0076
15.4682 | -34.8766 | -25.0156
15.4743 | -34.8771 | -25.0237
15.4804 | -34.8776 | -25.0317
15.4865 | -34.8781 | -25.0397
15.4926 | -34.8786 | -25.0477
15.4987 | -34.8791 | -25.056
15.5048 | -34.8796 | -25.0639
15.5109 | -34.8801 | -25.0722
15.517 | -34.8806 | -25.0803
15.5231 | -34.8811 | -25.0881
15.5292 | -34.8815 | -25.0962
15.5353 | -34.882 | -25.1043
15.5414 | -34.8825 | -25.1123
15.5476 | -34.883 | -25.1204
15.5537 | -34.8835 | -25.1286
15.5598 | -34.884 | -25.137
15.5659 | -34.8845 | -25.145
15.572 | -34.885 | -25.1531
15.5781 | -34.8855 | -25.1611
15.5842 | -34.886 | -25.1692
15.5903 | -34.8864 | -25.1771
15.5964 | -34.8869 | -25.1851
15.6025 | -34.8874 | -25.1932
15.6086 | -34.8879 | -25.2014
15.6147 | -34.8884 | -25.2095
15.6208 | -34.8889 | -25.2175
15.6269 | -34.8894 | -25.2254
15.633 | -34.8899 | -25.2335
15.6391 | -34.8904 | -25.2418
15.6452 | -34.8909 | -25.25
15.6513 | -34.8915 | -25.2577
15.6574 | -34.892 | -25.266
15.6635 | -34.8925 | -25.274
15.6696 | -34.8931 | -25.282
15.6757 | -34.8936 | -25.2903
15.6818 | -34.8941 | -25.2986
15.688 | -34.8946 | -25.3066
15.6941 | -34.8951 | -25.3146
15.7002 | -34.8956 | -25.3227
15.7063 | -34.8961 | -25.3307
15.7124 | -34.8966 | -25.3387
15.7185 | -34.8971 | -25.3466
15.7246 | -34.8976 | -25.3549
15.7307 | -34.8982 | -25.363
15.7368 | -34.8986 | -25.3708
15.7429 | -34.8992 | -25.3789
15.749 | -34.8997 | -25.3871
15.7551 | -34.9002 | -25.3951
15.7612 | -34.9007 | -25.4029
15.7673 | -34.9012 | -25.4108
15.7734 | -34.9017 | -25.4188
15.7795 | -34.9022 | -25.4269
15.7856 | -34.9027 | -25.4349
15.7917 | -34.9032 | -25.4428
15.7978 | -34.9037 | -25.4509
15.8039 | -34.9042 | -25.4588
15.81 | -34.9048 | -25.4666
15.8161 | -34.9052 | -25.4746
15.8222 | -34.9057 | -25.4825
15.8283 | -34.9063 | -25.4903
15.8345 | -34.9068 | -25.4981
15.8406 | -34.9073 | -25.5062
15.8467 | -34.9078 | -25.5142
15.8528 | -34.9083 | -25.5223
15.8589 | -34.9088 | -25.5303
15.865 | -34.9094 | -25.5384
15.8711 | -34.9099 | -25.5464
15.8772 | -34.9104 | -25.5543
15.8833 | -34.9109 | -25.5624
15.8894 | -34.9114 | -25.5705
15.8955 | -34.9119 | -25.5785
15.9016 | -34.9124 | -25.5864
15.9077 | -34.9129 | -25.5946
15.9138 | -34.9134 | -25.6028
15.9199 | -34.9139 | -25.6108
15.926 | -34.9144 | -25.6188
15.9321 | -34.9149 | -25.6266
15.9382 | -34.9154 | -25.6346
15.9443 | -34.9159 | -25.6426
15.9504 | -34.9164 | -25.6508
15.9565 | -34.917 | -25.6588
15.9626 | -34.9174 | -25.6668
15.9687 | -34.918 | -25.6746
15.9749 | -34.9185 | -25.6825
15.981 | -34.919 | -25.6904
15.9871 | -34.9195 | -25.6985
15.9932 | -34.92 | -25.7065
15.9993 | -34.9205 | -25.7144
16.0054 | -34.921 | -25.7224
16.0115 | -34.9215 | -25.7303
16.0176 | -34.922 | -25.7383
16.0237 | -34.9225 | -25.7462
16.0298 | -34.923 | -25.7541
16.0359 | -34.9235 | -25.762
16.042 | -34.9241 | -25.7701
16.0481 | -34.9246 | -25.778
16.0542 | -34.9251 | -25.786
16.0603 | -34.9256 | -25.7938
16.0664 | -34.9261 | -25.8019
16.0725 | -34.9266 | -25.8098
16.0786 | -34.9271 | -25.8177
16.0847 | -34.9276 | -25.8257
16.0908 | -34.9281 | -25.8336
16.0969 | -34.9286 | -25.8416
16.103 | -34.9291 | -25.8497
16.1091 | -34.9296 | -25.8578
16.1152 | -34.9301 | -25.8656
16.1214 | -34.9306 | -25.8735
16.1275 | -34.9311 | -25.8816
16.1336 | -34.9316 | -25.8894
16.1397 | -34.9321 | -25.8972
16.1458 | -34.9327 | -25.9051
16.1519 | -34.9332 | -25.9131
16.158 | -34.9337 | -25.9211
16.1641 | -34.9342 | -25.9291
16.1702 | -34.9347 | -25.9371
16.1763 | -34.9352 | -25.9449
16.1824 | -34.9357 | -25.9529
16.1885 | -34.9362 | -25.9606
16.1946 | -34.9368 | -25.9685
16.2007 | -34.9373 | -25.9765
16.2068 | -34.9378 | -25.9846
16.2129 | -34.9384 | -25.9926
16.219 | -34.9389 | -26.0001
16.2251 | -34.9394 | -26.0079
16.2312 | -34.9399 | -26.0158
16.2373 | -34.9404 | -26.0238
16.2434 | -34.9409 | -26.0318
16.2495 | -34.9414 | -26.0397
16.2556 | -34.9419 | -26.0476
16.2618 | -34.9425 | -26.0555
16.2679 | -34.943 | -26.0636
16.274 | -34.9435 | -26.0714
16.2801 | -34.9439 | -26.0794
16.2862 | -34.9445 | -26.0874
16.2923 | -34.945 | -26.0953
16.2984 | -34.9455 | -26.1033
16.3045 | -34.9461 | -26.1113
16.3106 | -34.9466 | -26.119
16.3167 | -34.9471 | -26.127
16.3228 | -34.9477 | -26.1346
16.3289 | -34.9481 | -26.1426
16.335 | -34.9487 | -26.1505
16.3411 | -34.9492 | -26.1583
16.3472 | -34.9497 | -26.1665
16.3533 | -34.9502 | -26.1744
16.3594 | -34.9507 | -26.182
16.3655 | -34.9512 | -26.19
16.3716 | -34.9517 | -26.198
16.3777 | -34.9522 | -26.2058
16.3838 | -34.9527 | -26.2138
16.3899 | -34.9532 | -26.2215
16.396 | -34.9537 | -26.2295
16.4021 | -34.9542 | -26.2373
16.4083 | -34.9547 | -26.245
16.4144 | -34.9552 | -26.2529
16.4205 | -34.9558 | -26.2608
16.4266 | -34.9562 | -26.2685
16.4327 | -34.9567 | -26.2763
16.4388 | -34.9572 | -26.2844
16.4449 | -34.9577 | -26.2922
16.451 | -34.9583 | -26.3002
16.4571 | -34.9588 | -26.3081
16.4632 | -34.9593 | -26.316
16.4693 | -34.9598 | -26.3238
16.4754 | -34.9603 | -26.3316
16.4815 | -34.9608 | -26.3395
16.4876 | -34.9613 | -26.3474
16.4937 | -34.9618 | -26.3552
16.4998 | -34.9624 | -26.3631
16.5059 | -34.9629 | -26.3711
16.512 | -34.9634 | -26.3791
16.5181 | -34.9639 | -26.3873
16.5242 | -34.9644 | -26.3954
16.5303 | -34.9649 | -26.4033
16.5364 | -34.9654 | -26.4112
16.5425 | -34.9659 | -26.4192
16.5487 | -34.9664 | -26.4271
16.5548 | -34.9669 | -26.4349
16.5609 | -34.9674 | -26.4427
16.567 | -34.9679 | -26.4507
16.5731 | -34.9684 | -26.4587
16.5792 | -34.9689 | -26.4665
16.5853 | -34.9695 | -26.4745
16.5914 | -34.9699 | -26.4825
16.5975 | -34.9705 | -26.4905
16.6036 | -34.971 | -26.4981
16.6097 | -34.9715 | -26.5063
16.6158 | -34.972 | -26.514
16.6219 | -34.9724 | -26.5219
16.628 | -34.973 | -26.5297
16.6341 | -34.9735 | -26.5375
16.6402 | -34.974 | -26.5454
16.6463 | -34.9745 | -26.5531
16.6524 | -34.975 | -26.5609
16.6585 | -34.9755 | -26.5691
16.6646 | -34.9761 | -26.5769
16.6707 | -34.9766 | -26.5848
16.6768 | -34.9771 | -26.5926
16.6829 | -34.9776 | -26.6003
16.689 | -34.9781 | -26.6082
16.6952 | -34.9786 | -26.6162
16.7013 | -34.9791 | -26.624
16.7074 | -34.9796 | -26.6317
16.7135 | -34.9801 | -26.6397
16.7196 | -34.9806 | -26.6475
16.7257 | -34.9812 | -26.6555
16.7318 | -34.9817 | -26.6636
16.7379 | -34.9822 | -26.6713
16.744 | -34.9827 | -26.6791
16.7501 | -34.9832 | -26.6869
16.7562 | -34.9837 | -26.6948
16.7623 | -34.9842 | -26.7025
16.7684 | -34.9847 | -26.7106
16.7745 | -34.9852 | -26.7185
16.7806 | -34.9857 | -26.7262
16.7867 | -34.9862 | -26.7342
16.7928 | -34.9867 | -26.7422
16.7989 | -34.9872 | -26.75
16.805 | -34.9877 | -26.7579
16.8111 | -34.9882 | -26.7657
16.8172 | -34.9888 | -26.7737
16.8233 | -34.9893 | -26.7815
16.8294 | -34.9898 | -26.7896
16.8356 | -34.9903 | -26.7975
16.8417 | -34.9908 | -26.8052
16.8478 | -34.9913 | -26.8128
16.8539 | -34.9918 | -26.8207
16.86 | -34.9923 | -26.8287
16.8661 | -34.9928 | -26.8364
16.8722 | -34.9933 | -26.8444
16.8783 | -34.9938 | -26.8523
16.8844 | -34.9943 | -26.8602
16.8905 | -34.9948 | -26.8681
16.8966 | -34.9953 | -26.8759
16.9027 | -34.9958 | -26.8838
16.9088 | -34.9963 | -26.8917
16.9149 | -34.9968 | -26.8996
16.921 | -34.9973 | -26.9071
16.9271 | -34.9978 | -26.915
16.9332 | -34.9983 | -26.923
16.9393 | -34.9988 | -26.9308
16.9454 | -34.9993 | -26.9386
16.9515 | -34.9998 | -26.9465
16.9576 | -35.0004 | -26.9542
16.9637 | -35.0009 | -26.962
16.9698 | -35.0014 | -26.9697
16.9759 | -35.0019 | -26.9775
16.9821 | -35.0024 | -26.9855
16.9882 | -35.0029 | -26.993
16.9943 | -35.0034 | -27.0008
17.0004 | -35.0039 | -27.0087
17.0065 | -35.0044 | -27.0163
17.0126 | -35.0049 | -27.0241
17.0187 | -35.0054 | -27.0318
17.0248 | -35.0059 | -27.0395
17.0309 | -35.0064 | -27.0473
17.037 | -35.0069 | -27.0551
17.0431 | -35.0074 | -27.0629
17.0492 | -35.0079 | -27.0707
17.0553 | -35.0084 | -27.0785
17.0614 | -35.0089 | -27.0863
17.0675 | -35.0094 | -27.0941
17.0736 | -35.0098 | -27.1021
17.0797 | -35.0103 | -27.1099
17.0858 | -35.0109 | -27.1179
17.0919 | -35.0114 | -27.1257
17.098 | -35.0119 | -27.1333
17.1041 | -35.0124 | -27.1411
17.1102 | -35.0129 | -27.1489
17.1163 | -35.0134 | -27.1566
17.1225 | -35.0139 | -27.1643
17.1286 | -35.0144 | -27.172
17.1347 | -35.015 | -27.1798
17.1408 | -35.0155 | -27.1877
17.1469 | -35.016 | -27.1952
17.153 | -35.0165 | -27.2028
17.1591 | -35.0169 | -27.2107
17.1652 | -35.0174 | -27.2185
17.1713 | -35.018 | -27.2264
17.1774 | -35.0184 | -27.2342
17.1835 | -35.0189 | -27.2419
17.1896 | -35.0194 | -27.2493
17.1957 | -35.0199 | -27.2571
17.2018 | -35.0204 | -27.2647
17.2079 | -35.021 | -27.2727
17.214 | -35.0215 | -27.2807
17.2201 | -35.022 | -27.2884
17.2262 | -35.0225 | -27.2964
17.2323 | -35.023 | -27.3041
17.2384 | -35.0235 | -27.3117
17.2445 | -35.024 | -27.3196
17.2506 | -35.0245 | -27.3273
17.2567 | -35.025 | -27.3351
17.2628 | -35.0255 | -27.3427
17.269 | -35.026 | -27.3503
17.2751 | -35.0265 | -27.3582
17.2812 | -35.0269 | -27.366
17.2873 | -35.0274 | -27.3738
17.2934 | -35.0279 | -27.3818
17.2995 | -35.0284 | -27.3898
17.3056 | -35.029 | -27.3974
17.3117 | -35.0295 | -27.4053
17.3178 | -35.03 | -27.4131
17.3239 | -35.0305 | -27.4208
17.33 | -35.031 | -27.4287
17.3361 | -35.0315 | -27.4364
17.3422 | -35.032 | -27.4439
17.3483 | -35.0325 | -27.4516
17.3544 | -35.033 | -27.4595
17.3605 | -35.0334 | -27.4672
17.3666 | -35.0339 | -27.4749
17.3727 | -35.0344 | -27.4825
17.3788 | -35.0349 | -27.49
17.3849 | -35.0354 | -27.4976
17.391 | -35.0359 | -27.5054
17.3971 | -35.0364 | -27.5131
17.4032 | -35.0369 | -27.5208
17.4094 | -35.0374 | -27.5282
17.4155 | -35.0379 | -27.5359
17.4216 | -35.0383 | -27.5435
17.4277 | -35.0388 | -27.5514
17.4338 | -35.0393 | -27.5591
17.4399 | -35.0398 | -27.567
17.446 | -35.0403 | -27.5745
17.4521 | -35.0409 | -27.5822
17.4582 | -35.0413 | -27.5898
17.4643 | -35.0418 | -27.5975
17.4704 | -35.0423 | -27.6052
17.4765 | -35.0428 | -27.6129
17.4826 | -35.0433 | -27.6206
17.4887 | -35.0438 | -27.6282
17.4948 | -35.0443 | -27.6358
17.5009 | -35.0448 | -27.6435
17.507 | -35.0453 | -27.6514
17.5131 | -35.0458 | -27.6592
17.5192 | -35.0462 | -27.6665
17.5253 | -35.0467 | -27.6743
17.5314 | -35.0472 | -27.6819
17.5375 | -35.0477 | -27.6896
17.5436 | -35.0482 | -27.6974
17.5497 | -35.0487 | -27.7052
17.5559 | -35.0492 | -27.7127
17.562 | -35.0496 | -27.7204
17.5681 | -35.0501 | -27.7279
17.5742 | -35.0506 | -27.7355
17.5803 | -35.0511 | -27.7432
17.5864 | -35.0516 | -27.751
17.5925 | -35.0521 | -27.7587
17.5986 | -35.0526 | -27.7664
17.6047 | -35.0531 | -27.7741
17.6108 | -35.0536 | -27.7817
17.6169 | -35.0541 | -27.7893
17.623 | -35.0546 | -27.7969
17.6291 | -35.0551 | -27.8046
17.6352 | -35.0556 | -27.8123
17.6413 | -35.0561 | -27.8201
17.6474 | -35.0565 | -27.8279
17.6535 | -35.057 | -27.8354
17.6596 | -35.0575 | -27.8428
17.6657 | -35.058 | -27.8504
17.6718 | -35.0585 | -27.8581
17.6779 | -35.0589 | -27.8656
17.684 | -35.0594 | -27.8732
17.6901 | -35.0599 | -27.8808
17.6963 | -35.0604 | -27.8883
17.7024 | -35.0609 | -27.8959
17.7085 | -35.0614 | -27.9037
17.7146 | -35.0619 | -27.9114
17.7207 | -35.0624 | -27.9188
17.7268 | -35.0628 | -27.9264
17.7329 | -35.0633 | -27.934
17.739 | -35.0638 | -27.9416
17.7451 | -35.0643 | -27.9491
17.7512 | -35.0648 | -27.9569
17.7573 | -35.0654 | -27.9649
17.7634 | -35.0658 | -27.9722
17.7695 | -35.0663 | -27.9796
17.7756 | -35.0668 | -27.9872
17.7817 | -35.0673 | -27.995
17.7878 | -35.0678 | -28.0026
17.7939 | -35.0683 | -28.0102
17.8 | -35.0688 | -28.0175
17.8061 | -35.0692 | -28.0251
17.8122 | -35.0697 | -28.0328
17.8183 | -35.0702 | -28.0405
17.8244 | -35.0707 | -28.048
17.8305 | -35.0712 | -28.0555
17.8366 | -35.0717 | -28.0626
17.8428 | -35.0722 | -28.0702
17.8489 | -35.0727 | -28.0775
17.855 | -35.0732 | -28.0849
17.8611 | -35.0737 | -28.0925
17.8672 | -35.0741 | -28.0998
17.8733 | -35.0746 | -28.1072
17.8794 | -35.0751 | -28.1149
17.8855 | -35.0756 | -28.1224
17.8916 | -35.0761 | -28.1302
17.8977 | -35.0765 | -28.1376
17.9038 | -35.077 | -28.145
17.9099 | -35.0775 | -28.1525
17.916 | -35.078 | -28.1602
17.9221 | -35.0785 | -28.1679
17.9282 | -35.079 | -28.1752
17.9343 | -35.0795 | -28.1829
17.9404 | -35.08 | -28.1905
17.9465 | -35.0805 | -28.198
17.9526 | -35.081 | -28.2055
17.9587 | -35.0815 | -28.2132
17.9648 | -35.082 | -28.2207
17.9709 | -35.0825 | -28.2281
17.977 | -35.0829 | -28.2354
17.9832 | -35.0834 | -28.243
17.9893 | -35.0839 | -28.2505
17.9954 | -35.0844 | -28.2579
18.0015 | -35.0849 | -28.2656
18.0076 | -35.0854 | -28.2728
18.0137 | -35.0859 | -28.2804
18.0198 | -35.0864 | -28.2879
18.0259 | -35.0868 | -28.2956
18.032 | -35.0873 | -28.303
18.0381 | -35.0878 | -28.3105
18.0442 | -35.0882 | -28.3179
18.0503 | -35.0887 | -28.3255
18.0564 | -35.0892 | -28.333
18.0625 | -35.0897 | -28.3408
18.0686 | -35.0901 | -28.3484
18.0747 | -35.0906 | -28.3561
18.0808 | -35.0911 | -28.3634
18.0869 | -35.0916 | -28.3708
18.093 | -35.0921 | -28.3783
18.0991 | -35.0926 | -28.3859
18.1052 | -35.0931 | -28.3935
18.1113 | -35.0936 | -28.401
18.1174 | -35.094 | -28.4082
18.1236 | -35.0945 | -28.4157
18.1297 | -35.095 | -28.4232
18.1358 | -35.0955 | -28.4306
18.1419 | -35.096 | -28.4383
18.148 | -35.0965 | -28.4457
18.1541 | -35.097 | -28.4529
18.1602 | -35.0974 | -28.4604
18.1663 | -35.0979 | -28.4676
18.1724 | -35.0984 | -28.475
18.1785 | -35.0989 | -28.4825
18.1846 | -35.0994 | -28.4898
18.1907 | -35.0998 | -28.4972
18.1968 | -35.1003 | -28.5047
18.2029 | -35.1008 | -28.512
18.209 | -35.1012 | -28.5196
18.2151 | -35.1017 | -28.5274
18.2212 | -35.1022 | -28.5346
18.2273 | -35.1027 | -28.5419
18.2334 | -35.1032 | -28.5495
18.2395 | -35.1037 | -28.557
18.2456 | -35.1041 | -28.5646
18.2517 | -35.1046 | -28.5718
18.2578 | -35.1051 | -28.579
18.2639 | -35.1056 | -28.5867
18.2701 | -35.1061 | -28.5943
18.2762 | -35.1066 | -28.6017
18.2823 | -35.1071 | -28.6089
18.2884 | -35.1076 | -28.6164
18.2945 | -35.1081 | -28.6235
18.3006 | -35.1086 | -28.6311
18.3067 | -35.1091 | -28.6385
18.3128 | -35.1096 | -28.6461
18.3189 | -35.1101 | -28.6536
18.325 | -35.1106 | -28.6608
18.3311 | -35.111 | -28.6682
18.3372 | -35.1115 | -28.6757
18.3433 | -35.112 | -28.683
18.3494 | -35.1125 | -28.6904
18.3555 | -35.113 | -28.698
18.3616 | -35.1135 | -28.7054
18.3677 | -35.1139 | -28.713
18.3738 | -35.1144 | -28.7201
18.3799 | -35.1149 | -28.7275
18.386 | -35.1154 | -28.7347
18.3921 | -35.1159 | -28.7422
18.3982 | -35.1163 | -28.7497
18.4043 | -35.1168 | -28.7573
18.4105 | -35.1173 | -28.7646
18.4166 | -35.1177 | -28.7718
18.4227 | -35.1182 | -28.7793
18.4288 | -35.1187 | -28.7866
18.4349 | -35.1192 | -28.7939
18.441 | -35.1196 | -28.8013
18.4471 | -35.1201 | -28.8085
18.4532 | -35.1206 | -28.8161
18.4593 | -35.121 | -28.8234
18.4654 | -35.1215 | -28.8305
18.4715 | -35.122 | -28.8381
18.4776 | -35.1225 | -28.8457
18.4837 | -35.1229 | -28.8529
18.4898 | -35.1234 | -28.8603
18.4959 | -35.1239 | -28.8678
18.502 | -35.1244 | -28.8749
18.5081 | -35.1249 | -28.8824
18.5142 | -35.1253 | -28.8896
18.5203 | -35.1258 | -28.8971
18.5264 | -35.1263 | -28.9041
18.5325 | -35.1268 | -28.9116
18.5386 | -35.1273 | -28.9189
18.5447 | -35.1278 | -28.9262
18.5508 | -35.1282 | -28.9335
18.557 | -35.1287 | -28.9409
18.5631 | -35.1292 | -28.9484
18.5692 | -35.1297 | -28.9559
18.5753 | -35.1302 | -28.9631
18.5814 | -35.1307 | -28.9703
18.5875 | -35.1311 | -28.9777
18.5936 | -35.1316 | -28.985
18.5997 | -35.132 | -28.9924
18.6058 | -35.1325 | -28.9995
18.6119 | -35.133 | -29.0068
18.618 | -35.1335 | -29.0142
18.6241 | -35.1339 | -29.0216
18.6302 | -35.1344 | -29.029
18.6363 | -35.1349 | -29.0364
18.6424 | -35.1354 | -29.0436
18.6485 | -35.1358 | -29.051
18.6546 | -35.1364 | -29.0584
18.6607 | -35.1368 | -29.0657
18.6668 | -35.1373 | -29.0732
18.6729 | -35.1377 | -29.0804
18.679 | -35.1382 | -29.0875
18.6851 | -35.1387 | -29.0947
18.6912 | -35.1392 | -29.1018
18.6974 | -35.1396 | -29.1092
18.7035 | -35.1401 | -29.1166
18.7096 | -35.1406 | -29.124
18.7157 | -35.141 | -29.1311
18.7218 | -35.1415 | -29.1383
18.7279 | -35.142 | -29.1458
18.734 | -35.1425 | -29.1531
18.7401 | -35.143 | -29.1604
18.7462 | -35.1435 | -29.1675
18.7523 | -35.144 | -29.1747
18.7584 | -35.1444 | -29.1821
18.7645 | -35.1449 | -29.1892
18.7706 | -35.1454 | -29.1964
18.7767 | -35.1459 | -29.2037
18.7828 | -35.1463 | -29.2106
18.7889 | -35.1468 | -29.2179
18.795 | -35.1473 | -29.2252
18.8011 | -35.1478 | -29.2325
18.8072 | -35.1483 | -29.2399
18.8133 | -35.1487 | -29.2473
18.8194 | -35.1492 | -29.254
18.8255 | -35.1497 | -29.2615
18.8316 | -35.1502 | -29.2687
18.8377 | -35.1507 | -29.2757
18.8439 | -35.1512 | -29.2829
18.85 | -35.1516 | -29.2901
18.8561 | -35.1521 | -29.2974
18.8622 | -35.1526 | -29.3048
18.8683 | -35.1531 | -29.3118
18.8744 | -35.1536 | -29.3189
18.8805 | -35.1541 | -29.3261
18.8866 | -35.1545 | -29.3334
18.8927 | -35.155 | -29.3405
18.8988 | -35.1554 | -29.3479
18.9049 | -35.1559 | -29.355
18.911 | -35.1564 | -29.3621
18.9171 | -35.1569 | -29.3694
18.9232 | -35.1573 | -29.3763
18.9293 | -35.1578 | -29.3836
18.9354 | -35.1583 | -29.3906
18.9415 | -35.1588 | -29.3979
18.9476 | -35.1592 | -29.4052
18.9537 | -35.1597 | -29.4125
18.9598 | -35.1602 | -29.4196
18.9659 | -35.1606 | -29.4269
18.972 | -35.1611 | -29.4339
18.9781 | -35.1616 | -29.4411
18.9843 | -35.162 | -29.4482
18.9904 | -35.1625 | -29.4555
18.9965 | -35.163 | -29.4624
19.0026 | -35.1635 | -29.4694
19.0087 | -35.164 | -29.4766
19.0148 | -35.1645 | -29.4837
19.0209 | -35.1649 | -29.4911
19.027 | -35.1654 | -29.4983
19.0331 | -35.1659 | -29.5051
19.0392 | -35.1664 | -29.5125
19.0453 | -35.1668 | -29.5197
19.0514 | -35.1673 | -29.5269
19.0575 | -35.1678 | -29.5341
19.0636 | -35.1683 | -29.5406
19.0697 | -35.1688 | -29.5477
19.0758 | -35.1692 | -29.555
19.0819 | -35.1697 | -29.5621
19.088 | -35.1702 | -29.5692
19.0941 | -35.1707 | -29.5765
19.1002 | -35.1712 | -29.5834
19.1063 | -35.1716 | -29.5905
19.1124 | -35.1721 | -29.5978
19.1185 | -35.1726 | -29.605
19.1246 | -35.1731 | -29.6123
19.1308 | -35.1735 | -29.6197
19.1369 | -35.174 | -29.6266
19.143 | -35.1745 | -29.6339
19.1491 | -35.1749 | -29.6411
19.1552 | -35.1753 | -29.6481
19.1613 | -35.1758 | -29.6551
19.1674 | -35.1763 | -29.6621
19.1735 | -35.1768 | -29.669
19.1796 | -35.1773 | -29.6762
19.1857 | -35.1778 | -29.6834
19.1918 | -35.1782 | -29.6905
19.1979 | -35.1787 | -29.6975
19.204 | -35.1792 | -29.7046
19.2101 | -35.1796 | -29.7117
19.2162 | -35.1801 | -29.7187
19.2223 | -35.1806 | -29.7259
19.2284 | -35.1811 | -29.7329
19.2345 | -35.1815 | -29.7399
19.2406 | -35.182 | -29.7467
19.2467 | -35.1825 | -29.754
19.2528 | -35.183 | -29.7611
19.2589 | -35.1834 | -29.7684
19.265 | -35.1839 | -29.7756
19.2712 | -35.1844 | -29.7828
19.2773 | -35.1848 | -29.7898
19.2834 | -35.1853 | -29.7968
19.2895 | -35.1858 | -29.804
19.2956 | -35.1863 | -29.8111
19.3017 | -35.1867 | -29.818
19.3078 | -35.1872 | -29.8252
19.3139 | -35.1877 | -29.832
19.32 | -35.1882 | -29.8391
19.3261 | -35.1887 | -29.8463
19.3322 | -35.1891 | -29.8532
19.3383 | -35.1896 | -29.8601
19.3444 | -35.1901 | -29.8667
19.3505 | -35.1906 | -29.8739
19.3566 | -35.1911 | -29.8808
19.3627 | -35.1915 | -29.8878
19.3688 | -35.192 | -29.8948
19.3749 | -35.1925 | -29.902
19.381 | -35.193 | -29.9089
19.3871 | -35.1935 | -29.9157
19.3932 | -35.1939 | -29.9228
19.3993 | -35.1944 | -29.9299
19.4054 | -35.1949 | -29.9371
19.4115 | -35.1954 | -29.9442
19.4177 | -35.1958 | -29.9511
19.4238 | -35.1963 | -29.9582
19.4299 | -35.1968 | -29.9653
19.436 | -35.1973 | -29.9723
19.4421 | -35.1977 | -29.9793
19.4482 | -35.1982 | -29.9864
19.4543 | -35.1988 | -29.993
19.4604 | -35.1992 | -29.9999
19.4665 | -35.1997 | -30.0068
19.4726 | -35.2001 | -30.0135
19.4787 | -35.2006 | -30.0204
19.4848 | -35.201 | -30.0275
19.4909 | -35.2015 | -30.0342
19.497 | -35.202 | -30.0413
19.5031 | -35.2024 | -30.0484
19.5092 | -35.2029 | -30.0558
19.5153 | -35.2033 | -30.0626
19.5214 | -35.2038 | -30.0694
19.5275 | -35.2042 | -30.0763
19.5336 | -35.2047 | -30.0832
19.5397 | -35.2052 | -30.0901
19.5458 | -35.2056 | -30.097
19.5519 | -35.206 | -30.104
19.5581 | -35.2065 | -30.1109
19.5642 | -35.2069 | -30.1178
19.5703 | -35.2074 | -30.125
19.5764 | -35.2079 | -30.1322
19.5825 | -35.2084 | -30.1394
19.5886 | -35.2088 | -30.1463
19.5947 | -35.2091 | -30.1541
19.6008 | -35.2096 | -30.1601
19.6069 | -35.2101 | -30.1667
19.613 | -35.2107 | -30.1742
19.6191 | -35.2111 | -30.1813
19.6252 | -35.2117 | -30.1882
19.6313 | -35.2121 | -30.1945
19.6374 | -35.2125 | -30.2016
19.6435 | -35.213 | -30.2086
19.6496 | -35.2135 | -30.2153
19.6557 | -35.2139 | -30.2224
19.6618 | -35.2144 | -30.2292
19.6679 | -35.2149 | -30.2361
19.674 | -35.2154 | -30.2431
19.6801 | -35.2159 | -30.2499
19.6862 | -35.2163 | -30.2571
19.6923 | -35.2168 | -30.2639
19.6984 | -35.2173 | -30.2703
19.7046 | -35.2178 | -30.2774
19.7107 | -35.2183 | -30.2842
19.7168 | -35.2188 | -30.2914
19.7229 | -35.2193 | -30.2987
19.729 | -35.2198 | -30.3057
19.7351 | -35.2202 | -30.3124
19.7412 | -35.2207 | -30.3195
19.7473 | -35.2211 | -30.3264
19.7534 | -35.2216 | -30.3332
19.7595 | -35.222 | -30.34
19.7656 | -35.2225 | -30.3471
19.7717 | -35.2229 | -30.354
19.7778 | -35.2234 | -30.3608
19.7839 | -35.2238 | -30.3677
19.79 | -35.2243 | -30.3742
19.7961 | -35.2247 | -30.3811
19.8022 | -35.2252 | -30.388
19.8083 | -35.2256 | -30.3948
19.8144 | -35.2261 | -30.4019
19.8205 | -35.2266 | -30.4091
19.8266 | -35.2271 | -30.4161
19.8327 | -35.2276 | -30.423
19.8388 | -35.228 | -30.4295
19.845 | -35.2285 | -30.4364
19.8511 | -35.2289 | -30.4435
19.8572 | -35.2293 | -30.4502
19.8633 | -35.2298 | -30.4569
19.8694 | -35.2303 | -30.4642
19.8755 | -35.2307 | -30.4709
19.8816 | -35.2312 | -30.4778
19.8877 | -35.2317 | -30.4848
19.8938 | -35.2321 | -30.4916
19.8999 | -35.2326 | -30.4987
19.906 | -35.233 | -30.5057
19.9121 | -35.2335 | -30.5124
19.9182 | -35.234 | -30.519
19.9243 | -35.2345 | -30.5257
19.9304 | -35.2349 | -30.5323
19.9365 | -35.2354 | -30.5395
19.9426 | -35.2358 | -30.546
19.9487 | -35.2363 | -30.5525
19.9548 | -35.2368 | -30.5599
19.9609 | -35.2372 | -30.5666
19.967 | -35.2377 | -30.5735
19.9731 | -35.2382 | -30.5805
19.9792 | -35.2386 | -30.5871
19.9853 | -35.2391 | -30.5942
19.9915 | -35.2395 | -30.6011
19.9976 | -35.24 | -30.608
20.0037 | -35.2404 | -30.6145
20.0098 | -35.2409 | -30.6216
20.0159 | -35.2413 | -30.6284
20.022 | -35.2418 | -30.6351
20.0281 | -35.2423 | -30.642
20.0342 | -35.2427 | -30.649
20.0403 | -35.2432 | -30.6558
20.0464 | -35.2436 | -30.6629
20.0525 | -35.2441 | -30.6696
20.0586 | -35.2446 | -30.6765
20.0647 | -35.245 | -30.6837
20.0708 | -35.2455 | -30.6907
20.0769 | -35.2459 | -30.6978
20.083 | -35.2465 | -30.7043
20.0891 | -35.2469 | -30.7115
20.0952 | -35.2473 | -30.7183
20.1013 | -35.2478 | -30.7253
20.1074 | -35.2483 | -30.732
20.1135 | -35.2487 | -30.739
20.1196 | -35.2491 | -30.7454
20.1257 | -35.2496 | -30.7523
20.1319 | -35.2501 | -30.7591
20.138 | -35.2505 | -30.7657
20.1441 | -35.251 | -30.7726
20.1502 | -35.2515 | -30.7796
20.1563 | -35.252 | -30.7861
20.1624 | -35.2524 | -30.7924
20.1685 | -35.2529 | -30.7994
20.1746 | -35.2533 | -30.8058
20.1807 | -35.2538 | -30.8125
20.1868 | -35.2543 | -30.8191
20.1929 | -35.2548 | -30.8256
20.199 | -35.2552 | -30.8324
20.2051 | -35.2557 | -30.8393
20.2112 | -35.2561 | -30.8462
20.2173 | -35.2566 | -30.8533
20.2234 | -35.2571 | -30.8602
20.2295 | -35.2576 | -30.8668
20.2356 | -35.258 | -30.8736
20.2417 | -35.2584 | -30.8807
20.2478 | -35.2589 | -30.8873
20.2539 | -35.2593 | -30.8939
20.26 | -35.2598 | -30.9003
20.2661 | -35.2603 | -30.907
20.2723 | -35.2607 | -30.9139
20.2784 | -35.2611 | -30.9209
20.2845 | -35.2615 | -30.9278
20.2906 | -35.262 | -30.9343
20.2967 | -35.2624 | -30.9404
20.3028 | -35.2629 | -30.9471
20.3089 | -35.2634 | -30.9535
20.315 | -35.2639 | -30.9602
20.3211 | -35.2642 | -30.9668
20.3272 | -35.2647 | -30.9736
20.3333 | -35.2652 | -30.98
20.3394 | -35.2656 | -30.9867
20.3455 | -35.2661 | -30.9934
20.3516 | -35.2666 | -31.0004
20.3577 | -35.267 | -31.0073
20.3638 | -35.2675 | -31.0142
20.3699 | -35.2679 | -31.0209
20.376 | -35.2684 | -31.0275
20.3821 | -35.2688 | -31.0347
20.3882 | -35.2692 | -31.0411
20.3943 | -35.2697 | -31.0478
20.4004 | -35.2702 | -31.0542
20.4065 | -35.2706 | -31.0608
20.4126 | -35.2711 | -31.0675
20.4188 | -35.2715 | -31.0743
20.4249 | -35.272 | -31.0811
20.431 | -35.2724 | -31.088
20.4371 | -35.2729 | -31.0948
20.4432 | -35.2733 | -31.1013
20.4493 | -35.2737 | -31.1078
20.4554 | -35.2742 | -31.1148
20.4615 | -35.2746 | -31.1213
20.4676 | -35.2751 | -31.1279
20.4737 | -35.2755 | -31.1343
20.4798 | -35.276 | -31.141
20.4859 | -35.2764 | -31.1481
20.492 | -35.2769 | -31.1545
20.4981 | -35.2774 | -31.1608
20.5042 | -35.2778 | -31.1675
20.5103 | -35.2783 | -31.1739
20.5164 | -35.2787 | -31.1805
20.5225 | -35.2792 | -31.1871
20.5286 | -35.2797 | -31.194
20.5347 | -35.2801 | -31.2011
20.5408 | -35.2806 | -31.2077
20.5469 | -35.281 | -31.2141
20.553 | -35.2815 | -31.2209
20.5592 | -35.282 | -31.228
20.5653 | -35.2825 | -31.2344
20.5714 | -35.2829 | -31.2411
20.5775 | -35.2835 | -31.2474
20.5836 | -35.2839 | -31.254
20.5897 | -35.2844 | -31.261
20.5958 | -35.2849 | -31.2675
20.6019 | -35.2853 | -31.2743
20.608 | -35.2857 | -31.2807
20.6141 | -35.2861 | -31.2873
20.6202 | -35.2866 | -31.2939
20.6263 | -35.2871 | -31.3002
20.6324 | -35.2876 | -31.3065
20.6385 | -35.2881 | -31.3131
20.6446 | -35.2885 | -31.3201
20.6507 | -35.289 | -31.327
20.6568 | -35.2895 | -31.3334
20.6629 | -35.29 | -31.3399
20.669 | -35.2905 | -31.3468
20.6751 | -35.2909 | -31.3537
20.6812 | -35.2915 | -31.3602
20.6873 | -35.2919 | -31.3665
20.6934 | -35.2924 | -31.3729
20.6995 | -35.2929 | -31.3797
20.7057 | -35.2933 | -31.3866
20.7118 | -35.2939 | -31.3931
20.7179 | -35.2943 | -31.3989
20.724 | -35.2947 | -31.4057
20.7301 | -35.2951 | -31.4124
20.7362 | -35.2956 | -31.4188
20.7423 | -35.296 | -31.4256
20.7484 | -35.2964 | -31.4318
20.7545 | -35.2969 | -31.4383
20.7606 | -35.2973 | -31.445
20.7667 | -35.2978 | -31.4519
20.7728 | -35.2982 | -31.4589
20.7789 | -35.2986 | -31.466
20.785 | -35.2991 | -31.473
20.7911 | -35.2995 | -31.4798
20.7972 | -35.3 | -31.4867
20.8033 | -35.3005 | -31.4936
20.8094 | -35.3009 | -31.4998
20.8155 | -35.3014 | -31.506
20.8216 | -35.3019 | -31.5125
20.8277 | -35.3023 | -31.5191
20.8338 | -35.3027 | -31.5257
20.8399 | -35.3032 | -31.5326
20.8461 | -35.3036 | -31.5393
20.8522 | -35.3041 | -31.5461
20.8583 | -35.3046 | -31.5532
20.8644 | -35.3051 | -31.5593
20.8705 | -35.3056 | -31.5659
20.8766 | -35.306 | -31.5725
20.8827 | -35.3065 | -31.5791
20.8888 | -35.3069 | -31.5858
20.8949 | -35.3073 | -31.5921
20.901 | -35.3078 | -31.5984
20.9071 | -35.3082 | -31.6052
20.9132 | -35.3086 | -31.6121
20.9193 | -35.309 | -31.6186
20.9254 | -35.3095 | -31.6254
20.9315 | -35.3099 | -31.6324
20.9376 | -35.3103 | -31.6386
20.9437 | -35.3108 | -31.6453
20.9498 | -35.3112 | -31.6515
20.9559 | -35.3117 | -31.6582
20.962 | -35.3122 | -31.6649
20.9681 | -35.3127 | -31.6713
20.9742 | -35.3131 | -31.6777
20.9803 | -35.3136 | -31.6845
20.9864 | -35.314 | -31.6908
20.9926 | -35.3144 | -31.6973
20.9987 | -35.3147 | -31.7032
21.0048 | -35.3152 | -31.7102
21.0109 | -35.3157 | -31.7169
21.017 | -35.3162 | -31.7236
21.0231 | -35.3167 | -31.7299
21.0292 | -35.3171 | -31.7361
21.0353 | -35.3175 | -31.7429
21.0414 | -35.318 | -31.7495
21.0475 | -35.3184 | -31.756
21.0536 | -35.3189 | -31.7623
21.0597 | -35.3193 | -31.7692
21.0658 | -35.3197 | -31.7758
21.0719 | -35.3201 | -31.782
21.078 | -35.3205 | -31.7885
21.0841 | -35.3211 | -31.7951
21.0902 | -35.3215 | -31.8017
21.0963 | -35.3219 | -31.8082
21.1024 | -35.3224 | -31.8146
21.1085 | -35.3229 | -31.8211
21.1146 | -35.3233 | -31.8278
21.1207 | -35.3237 | -31.834
21.1268 | -35.3242 | -31.8405
21.133 | -35.3246 | -31.8469
21.1391 | -35.3251 | -31.8537
21.1452 | -35.3255 | -31.86
21.1513 | -35.326 | -31.8667
21.1574 | -35.3265 | -31.8732
21.1635 | -35.327 | -31.8799
21.1696 | -35.3274 | -31.8861
21.1757 | -35.3278 | -31.8927
21.1818 | -35.3282 | -31.8989
21.1879 | -35.3287 | -31.9054
21.194 | -35.3291 | -31.9121
21.2001 | -35.3295 | -31.9185
21.2062 | -35.33 | -31.9253
21.2123 | -35.3305 | -31.932
21.2184 | -35.331 | -31.9387
21.2245 | -35.3314 | -31.9454
21.2306 | -35.3319 | -31.9523
21.2367 | -35.3323 | -31.9591
21.2428 | -35.3327 | -31.9655
21.2489 | -35.3332 | -31.9719
21.255 | -35.3336 | -31.9787
21.2611 | -35.3341 | -31.9855
21.2672 | -35.3346 | -31.9922
21.2733 | -35.3351 | -31.9992
21.2795 | -35.3356 | -32.0062
21.2856 | -35.336 | -32.0127
21.2917 | -35.3364 | -32.0193
21.2978 | -35.3368 | -32.0252
21.3039 | -35.3373 | -32.0323
21.31 | -35.3377 | -32.039
21.3161 | -35.3381 | -32.0446
21.3222 | -35.3386 | -32.0514
21.3283 | -35.3391 | -32.0575
21.3344 | -35.3395 | -32.0634
21.3405 | -35.34 | -32.0699
21.3466 | -35.3404 | -32.0759
21.3527 | -35.3408 | -32.082
21.3588 | -35.3413 | -32.0888
21.3649 | -35.3417 | -32.0955
21.371 | -35.3421 | -32.1021
21.3771 | -35.3426 | -32.1087
21.3832 | -35.343 | -32.1148
21.3893 | -35.3434 | -32.1208
21.3954 | -35.3438 | -32.1276
21.4015 | -35.3442 | -32.1338
21.4076 | -35.3447 | -32.1407
21.4137 | -35.3451 | -32.1475
21.4199 | -35.3456 | -32.1535
21.426 | -35.346 | -32.16
21.4321 | -35.3465 | -32.1669
21.4382 | -35.3469 | -32.1732
21.4443 | -35.3474 | -32.1801
21.4504 | -35.3479 | -32.1858
21.4565 | -35.3484 | -32.1918
21.4626 | -35.3489 | -32.1977
21.4687 | -35.3493 | -32.2045
21.4748 | -35.3499 | -32.2107
21.4809 | -35.3503 | -32.217
21.487 | -35.3506 | -32.2233
21.4931 | -35.3511 | -32.2295
21.4992 | -35.3516 | -32.2356
21.5053 | -35.352 | -32.2422
21.5114 | -35.3524 | -32.2485
21.5175 | -35.3529 | -32.2549
21.5236 | -35.3533 | -32.2614
21.5297 | -35.3537 | -32.2676
21.5358 | -35.3542 | -32.274
21.5419 | -35.3546 | -32.2802
21.548 | -35.355 | -32.2869
21.5541 | -35.3555 | -32.2937
21.5602 | -35.3559 | -32.3007
21.5664 | -35.3564 | -32.3065
21.5725 | -35.3568 | -32.3133
21.5786 | -35.3573 | -32.3201
21.5847 | -35.3577 | -32.3272
21.5908 | -35.3581 | -32.3341
21.5969 | -35.3586 | -32.3404
21.603 | -35.359 | -32.3468
21.6091 | -35.3594 | -32.353
21.6152 | -35.3599 | -32.3591
21.6213 | -35.3602 | -32.3657
21.6274 | -35.3608 | -32.3723
21.6335 | -35.3611 | -32.3782
21.6396 | -35.3615 | -32.3849
21.6457 | -35.362 | -32.3914
21.6518 | -35.3624 | -32.398
21.6579 | -35.3628 | -32.4048
21.664 | -35.3633 | -32.4111
21.6701 | -35.3638 | -32.4173
21.6762 | -35.3642 | -32.4239
21.6823 | -35.3646 | -32.43
21.6884 | -35.365 | -32.4368
21.6945 | -35.3654 | -32.443
21.7006 | -35.3658 | -32.4498
21.7068 | -35.3662 | -32.4556
21.7129 | -35.3667 | -32.462
21.719 | -35.3672 | -32.4683
21.7251 | -35.3676 | -32.4754
21.7312 | -35.368 | -32.4824
21.7373 | -35.3685 | -32.4886
21.7434 | -35.369 | -32.4949
21.7495 | -35.3694 | -32.5009
21.7556 | -35.3698 | -32.5075
21.7617 | -35.3702 | -32.5139
21.7678 | -35.3707 | -32.5202
21.7739 | -35.3712 | -32.5265
21.78 | -35.3717 | -32.5335
21.7861 | -35.3721 | -32.5404
21.7922 | -35.3726 | -32.5474
21.7983 | -35.373 | -32.5536
21.8044 | -35.3734 | -32.5596
21.8105 | -35.3739 | -32.5657
21.8166 | -35.3743 | -32.5721
21.8227 | -35.3748 | -32.5781
21.8288 | -35.3753 | -32.5843
21.8349 | -35.3757 | -32.5905
21.841 | -35.3762 | -32.5972
21.8471 | -35.3767 | -32.6034
21.8533 | -35.3771 | -32.6096
21.8594 | -35.3775 | -32.6155
21.8655 | -35.378 | -32.6218
21.8716 | -35.3783 | -32.628
21.8777 | -35.3789 | -32.6344
21.8838 | -35.3794 | -32.6406
21.8899 | -35.3797 | -32.647
21.896 | -35.3802 | -32.6529
21.9021 | -35.3806 | -32.6589
21.9082 | -35.3811 | -32.6649
21.9143 | -35.3815 | -32.6712
21.9204 | -35.3821 | -32.6775
21.9265 | -35.3826 | -32.6843
21.9326 | -35.3829 | -32.6906
21.9387 | -35.3834 | -32.6976
21.9448 | -35.3838 | -32.7044
21.9509 | -35.3842 | -32.711
21.957 | -35.3847 | -32.717
21.9631 | -35.3851 | -32.7238
21.9692 | -35.3856 | -32.7306
21.9753 | -35.3861 | -32.7365
21.9814 | -35.3865 | -32.7427
21.9875 | -35.3869 | -32.7495
21.9937 | -35.3873 | -32.7554
21.9998 | -35.3877 | -32.7619
22.0059 | -35.388 | -32.7682
22.012 | -35.3885 | -32.775
22.0181 | -35.3889 | -32.7811
22.0242 | -35.3895 | -32.7872
22.0303 | -35.3899 | -32.7939
22.0364 | -35.3903 | -32.8001
22.0425 | -35.3908 | -32.8058
22.0486 | -35.3912 | -32.8127
22.0547 | -35.3917 | -32.8186
22.0608 | -35.3922 | -32.8254
22.0669 | -35.3927 | -32.8318
22.073 | -35.3932 | -32.8384
22.0791 | -35.3936 | -32.8447
22.0852 | -35.394 | -32.8507
22.0913 | -35.3944 | -32.8571
22.0974 | -35.3948 | -32.8635
22.1035 | -35.3952 | -32.8698
22.1096 | -35.3956 | -32.8764
22.1157 | -35.3961 | -32.8829
22.1218 | -35.3965 | -32.8889
22.1279 | -35.3969 | -32.8946
22.134 | -35.3973 | -32.9009
22.1402 | -35.3978 | -32.9075
22.1463 | -35.3983 | -32.9138
22.1524 | -35.3988 | -32.9207
22.1585 | -35.3992 | -32.9272
22.1646 | -35.3997 | -32.9328
22.1707 | -35.4002 | -32.9395
22.1768 | -35.4007 | -32.9459
22.1829 | -35.4011 | -32.9522
22.189 | -35.4016 | -32.9587
22.1951 | -35.4021 | -32.965
22.2012 | -35.4025 | -32.9719
22.2073 | -35.4029 | -32.9781
22.2134 | -35.4033 | -32.9846
22.2195 | -35.4038 | -32.991
22.2256 | -35.4042 | -32.9971
22.2317 | -35.4046 | -33.0031
22.2378 | -35.405 | -33.0095
22.2439 | -35.4054 | -33.0164
22.25 | -35.4057 | -33.0225
22.2561 | -35.4062 | -33.0281
22.2622 | -35.4067 | -33.0348
22.2683 | -35.4071 | -33.0417
22.2744 | -35.4076 | -33.0478
22.2806 | -35.4081 | -33.054
22.2867 | -35.4085 | -33.0605
22.2928 | -35.4089 | -33.0674
22.2989 | -35.4093 | -33.0737
22.305 | -35.4096 | -33.0803
22.3111 | -35.41 | -33.087
22.3172 | -35.4104 | -33.0937
22.3233 | -35.4109 | -33.1004
22.3294 | -35.4113 | -33.1074
22.3355 | -35.4117 | -33.1144
22.3416 | -35.4122 | -33.1206
22.3477 | -35.4127 | -33.1266
22.3538 | -35.4132 | -33.1333
22.3599 | -35.4136 | -33.1403
22.366 | -35.414 | -33.1469
22.3721 | -35.4144 | -33.1535
22.3782 | -35.4148 | -33.1603
22.3843 | -35.4152 | -33.1664
22.3904 | -35.4157 | -33.1728
22.3965 | -35.416 | -33.1798
22.4026 | -35.4166 | -33.1866
22.4087 | -35.417 | -33.1921
22.4148 | -35.4175 | -33.198
22.4209 | -35.4179 | -33.2048
22.4271 | -35.4183 | -33.2112
22.4332 | -35.4187 | -33.2175
22.4393 | -35.4191 | -33.2234
22.4454 | -35.4196 | -33.2302
22.4515 | -35.42 | -33.2362
22.4576 | -35.4205 | -33.2428
22.4637 | -35.4209 | -33.2492
22.4698 | -35.4213 | -33.2557
22.4759 | -35.4217 | -33.2622
22.482 | -35.4222 | -33.2689
22.4881 | -35.4225 | -33.2758
22.4942 | -35.423 | -33.2826
22.5003 | -35.4234 | -33.2892
22.5064 | -35.4239 | -33.2953
22.5125 | -35.4243 | -33.3018
22.5186 | -35.4247 | -33.3082
22.5247 | -35.4251 | -33.3146
22.5308 | -35.4256 | -33.3209
22.5369 | -35.426 | -33.3271
22.543 | -35.4265 | -33.333
22.5491 | -35.4269 | -33.3392
22.5552 | -35.4273 | -33.3457
22.5613 | -35.4277 | -33.3523
22.5675 | -35.4282 | -33.3587
22.5736 | -35.4287 | -33.3655
22.5797 | -35.4291 | -33.372
22.5858 | -35.4295 | -33.3781
22.5919 | -35.43 | -33.3843
22.598 | -35.4304 | -33.3904
22.6041 | -35.4309 | -33.3966
22.6102 | -35.4312 | -33.4031
22.6163 | -35.4317 | -33.4093
22.6224 | -35.4322 | -33.4155
22.6285 | -35.4326 | -33.4214
22.6346 | -35.433 | -33.4277
22.6407 | -35.4334 | -33.4342
22.6468 | -35.4339 | -33.4404
22.6529 | -35.4343 | -33.4468
22.659 | -35.4347 | -33.4529
22.6651 | -35.4351 | -33.4588
22.6712 | -35.4356 | -33.4652
22.6773 | -35.436 | -33.4715
22.6834 | -35.4365 | -33.4778
22.6895 | -35.4369 | -33.4838
22.6956 | -35.4373 | -33.4899
22.7017 | -35.4378 | -33.4964
22.7079 | -35.4383 | -33.5025
22.714 | -35.4388 | -33.509
22.7201 | -35.4393 | -33.5152
22.7262 | -35.4396 | -33.5225
22.7323 | -35.44 | -33.5284
22.7384 | -35.4404 | -33.5349
22.7445 | -35.4409 | -33.541
22.7506 | -35.4414 | -33.5478
22.7567 | -35.4418 | -33.5543
22.7628 | -35.4422 | -33.561
22.7689 | -35.4427 | -33.5669
22.775 | -35.4431 | -33.5738
22.7811 | -35.4435 | -33.5801
22.7872 | -35.444 | -33.5866
22.7933 | -35.4444 | -33.592
22.7994 | -35.4449 | -33.5982
22.8055 | -35.4453 | -33.6042
22.8116 | -35.4457 | -33.6103
22.8177 | -35.4461 | -33.6164
22.8238 | -35.4466 | -33.6226
22.8299 | -35.4469 | -33.6285
22.836 | -35.4474 | -33.6346
22.8421 | -35.4478 | -33.6408
22.8482 | -35.4483 | -33.6469
22.8544 | -35.4487 | -33.6535
22.8605 | -35.4492 | -33.6601
22.8666 | -35.4495 | -33.666
22.8727 | -35.45 | -33.672
22.8788 | -35.4504 | -33.6781
22.8849 | -35.4508 | -33.6842
22.891 | -35.4512 | -33.6905
22.8971 | -35.4517 | -33.6966
22.9032 | -35.4522 | -33.7027
22.9093 | -35.4527 | -33.7092
22.9154 | -35.4531 | -33.7155
22.9215 | -35.4536 | -33.7218
22.9276 | -35.454 | -33.728
22.9337 | -35.4545 | -33.7335
22.9398 | -35.4549 | -33.7398
22.9459 | -35.4553 | -33.7463
22.952 | -35.4557 | -33.7522
22.9581 | -35.4562 | -33.7588
22.9642 | -35.4566 | -33.765
22.9703 | -35.457 | -33.7711
22.9764 | -35.4575 | -33.7772
22.9825 | -35.4579 | -33.7836
22.9886 | -35.4583 | -33.79
22.9948 | -35.4587 | -33.7961
23.0009 | -35.4591 | -33.8022
23.007 | -35.4596 | -33.8085
23.0131 | -35.4601 | -33.8146
23.0192 | -35.4605 | -33.8208
23.0253 | -35.4609 | -33.8272
23.0314 | -35.4613 | -33.8337
23.0375 | -35.4618 | -33.8401
23.0436 | -35.4622 | -33.8462
23.0497 | -35.4626 | -33.8526
23.0558 | -35.4631 | -33.8592
23.0619 | -35.4635 | -33.8656
23.068 | -35.4639 | -33.872
23.0741 | -35.4644 | -33.8783
23.0802 | -35.4649 | -33.8841
23.0863 | -35.4653 | -33.8903
23.0924 | -35.4657 | -33.8961
23.0985 | -35.4662 | -33.9025
23.1046 | -35.4665 | -33.9085
23.1107 | -35.467 | -33.915
23.1168 | -35.4675 | -33.9214
23.1229 | -35.4679 | -33.9278
23.129 | -35.4683 | -33.9342
23.1351 | -35.4688 | -33.9404
23.1413 | -35.4693 | -33.9465
23.1474 | -35.4697 | -33.9524
23.1535 | -35.4702 | -33.9588
23.1596 | -35.4707 | -33.9652
23.1657 | -35.4711 | -33.9713
23.1718 | -35.4716 | -33.9778
23.1779 | -35.4721 | -33.9842
23.184 | -35.4724 | -33.9901
23.1901 | -35.4729 | -33.9962
23.1962 | -35.4733 | -34.0024
23.2023 | -35.4738 | -34.0089
23.2084 | -35.4742 | -34.0152
23.2145 | -35.4746 | -34.0216
23.2206 | -35.475 | -34.028
23.2267 | -35.4754 | -34.0339
23.2328 | -35.4758 | -34.0404
23.2389 | -35.4763 | -34.0466
23.245 | -35.4768 | -34.053
23.2511 | -35.4773 | -34.059
23.2572 | -35.4777 | -34.0652
23.2633 | -35.4782 | -34.071
23.2694 | -35.4786 | -34.0776
23.2755 | -35.4791 | -34.0841
23.2817 | -35.4795 | -34.0901
23.2878 | -35.48 | -34.0963
23.2939 | -35.4805 | -34.1024
23.3 | -35.4809 | -34.1082
23.3061 | -35.4813 | -34.1146
23.3122 | -35.4817 | -34.121
23.3183 | -35.4821 | -34.1275
23.3244 | -35.4824 | -34.1334
23.3305 | -35.4828 | -34.1393
23.3366 | -35.4832 | -34.1454
23.3427 | -35.4837 | -34.1513
23.3488 | -35.4841 | -34.1577
23.3549 | -35.4846 | -34.1642
23.361 | -35.4851 | -34.1699
23.3671 | -35.4855 | -34.1761
23.3732 | -35.4859 | -34.1825
23.3793 | -35.4863 | -34.1889
23.3854 | -35.4867 | -34.1949
23.3915 | -35.4871 | -34.2008
23.3976 | -35.4876 | -34.2073
23.4037 | -35.488 | -34.2133
23.4098 | -35.4885 | -34.2197
23.4159 | -35.4889 | -34.2264
23.422 | -35.4894 | -34.2322
23.4282 | -35.4898 | -34.238
23.4343 | -35.4902 | -34.2441
23.4404 | -35.4907 | -34.2502
23.4465 | -35.4912 | -34.2567
23.4526 | -35.4916 | -34.2627
23.4587 | -35.492 | -34.2686
23.4648 | -35.4925 | -34.2745
23.4709 | -35.4929 | -34.2802
23.477 | -35.4933 | -34.2863
23.4831 | -35.4937 | -34.292
23.4892 | -35.4942 | -34.2979
23.4953 | -35.4946 | -34.3038
23.5014 | -35.495 | -34.31
23.5075 | -35.4954 | -34.3163
23.5136 | -35.4958 | -34.3225
23.5197 | -35.4963 | -34.3289
23.5258 | -35.4967 | -34.3353
23.5319 | -35.4972 | -34.3413
23.538 | -35.4977 | -34.3475
23.5441 | -35.4981 | -34.3534
23.5502 | -35.4985 | -34.3597
23.5563 | -35.4989 | -34.3658
23.5624 | -35.4994 | -34.3723
23.5686 | -35.4999 | -34.3779
23.5747 | -35.5003 | -34.3839
23.5808 | -35.5007 | -34.3901
23.5869 | -35.501 | -34.396
23.593 | -35.5014 | -34.4023
23.5991 | -35.5019 | -34.4079
23.6052 | -35.5025 | -34.4145
23.6113 | -35.5029 | -34.4205
23.6174 | -35.5033 | -34.4266
23.6235 | -35.5037 | -34.4325
23.6296 | -35.5041 | -34.4381
23.6357 | -35.5046 | -34.4441
23.6418 | -35.5051 | -34.4508
23.6479 | -35.5055 | -34.4565
23.654 | -35.5059 | -34.4627
23.6601 | -35.5063 | -34.4685
23.6662 | -35.5067 | -34.4742
23.6723 | -35.5071 | -34.4795
23.6784 | -35.5075 | -34.4857
23.6845 | -35.5079 | -34.4921
23.6906 | -35.5084 | -34.4986
23.6967 | -35.5087 | -34.5053
23.7028 | -35.5092 | -34.5116
23.7089 | -35.5097 | -34.5169
23.7151 | -35.51 | -34.5228
23.7212 | -35.5104 | -34.529
23.7273 | -35.5107 | -34.5353
23.7334 | -35.5112 | -34.5411
23.7395 | -35.5117 | -34.5472
23.7456 | -35.512 | -34.5533
23.7517 | -35.5125 | -34.5595
23.7578 | -35.5129 | -34.5663
23.7639 | -35.5133 | -34.5725
23.77 | -35.5137 | -34.5785
23.7761 | -35.5141 | -34.5848
23.7822 | -35.5145 | -34.5899
23.7883 | -35.515 | -34.5959
23.7944 | -35.5154 | -34.6021
23.8005 | -35.5158 | -34.6087
23.8066 | -35.5163 | -34.6149
23.8127 | -35.5168 | -34.6217
23.8188 | -35.5173 | -34.628
23.8249 | -35.5177 | -34.634
23.831 | -35.5182 | -34.6401
23.8371 | -35.5186 | -34.6464
23.8432 | -35.519 | -34.6528
23.8493 | -35.5194 | -34.6594
23.8555 | -35.5198 | -34.6652
23.8616 | -35.5203 | -34.6712
23.8677 | -35.5207 | -34.6772
23.8738 | -35.5211 | -34.6831
23.8799 | -35.5216 | -34.6895
23.886 | -35.522 | -34.6951
23.8921 | -35.5224 | -34.7014
23.8982 | -35.5229 | -34.7073
23.9043 | -35.5233 | -34.7132
23.9104 | -35.5238 | -34.7191
23.9165 | -35.5242 | -34.7251
23.9226 | -35.5246 | -34.7312
23.9287 | -35.525 | -34.7371
23.9348 | -35.5254 | -34.7431
23.9409 | -35.5258 | -34.7492
23.947 | -35.5262 | -34.7552
23.9531 | -35.5267 | -34.761
23.9592 | -35.5271 | -34.7671
23.9653 | -35.5276 | -34.7736
23.9714 | -35.5279 | -34.7799
23.9775 | -35.5284 | -34.786
23.9836 | -35.5288 | -34.7917
23.9897 | -35.5293 | -34.7972
23.9958 | -35.5297 | -34.803
24.002 | -35.5302 | -34.8089
24.0081 | -35.5306 | -34.8151
24.0142 | -35.531 | -34.8215
24.0203 | -35.5315 | -34.8269
24.0264 | -35.5318 | -34.8333
24.0325 | -35.5323 | -34.8389
24.0386 | -35.5327 | -34.8453
24.0447 | -35.5332 | -34.8517
24.0508 | -35.5335 | -34.8576
24.0569 | -35.534 | -34.8633
24.063 | -35.5343 | -34.8692
24.0691 | -35.5348 | -34.875
24.0752 | -35.5351 | -34.8809
24.0813 | -35.5356 | -34.8873
24.0874 | -35.5361 | -34.8935
24.0935 | -35.5365 | -34.8996
24.0996 | -35.537 | -34.905
24.1057 | -35.5373 | -34.9108
24.1118 | -35.5378 | -34.9172
24.1179 | -35.5383 | -34.9233
24.124 | -35.5387 | -34.9299
24.1301 | -35.5391 | -34.9359
24.1362 | -35.5395 | -34.942
24.1424 | -35.5399 | -34.9482
24.1485 | -35.5403 | -34.9546
24.1546 | -35.5407 | -34.9609
24.1607 | -35.5411 | -34.9669
24.1668 | -35.5415 | -34.972
24.1729 | -35.5419 | -34.9784
24.179 | -35.5424 | -34.9842
24.1851 | -35.5427 | -34.9903
24.1912 | -35.5432 | -34.9961
24.1973 | -35.5435 | -35.0021
24.2034 | -35.544 | -35.0077
24.2095 | -35.5444 | -35.0135
24.2156 | -35.5449 | -35.0194
24.2217 | -35.5453 | -35.0256
24.2278 | -35.5458 | -35.0318
24.2339 | -35.5462 | -35.0379
24.24 | -35.5466 | -35.0436
24.2461 | -35.5471 | -35.0497
24.2522 | -35.5476 | -35.0556
24.2583 | -35.548 | -35.0615
24.2644 | -35.5484 | -35.0678
24.2705 | -35.5489 | -35.0741
24.2766 | -35.5492 | -35.0801
24.2827 | -35.5497 | -35.0856
24.2889 | -35.5501 | -35.0917
24.295 | -35.5505 | -35.0979
24.3011 | -35.5508 | -35.104
24.3072 | -35.5513 | -35.1098
24.3133 | -35.5517 | -35.1163
24.3194 | -35.5521 | -35.1221
24.3255 | -35.5525 | -35.1278
24.3316 | -35.5529 | -35.1335
24.3377 | -35.5534 | -35.1393
24.3438 | -35.5537 | -35.1448
24.3499 | -35.5541 | -35.1507
24.356 | -35.5545 | -35.1572
24.3621 | -35.5549 | -35.1632
24.3682 | -35.5553 | -35.1692
24.3743 | -35.5557 | -35.1751
24.3804 | -35.5561 | -35.1814
24.3865 | -35.5565 | -35.1868
24.3926 | -35.557 | -35.1926
24.3987 | -35.5573 | -35.1987
24.4048 | -35.5578 | -35.2045
24.4109 | -35.5582 | -35.21
24.417 | -35.5587 | -35.2157
24.4231 | -35.5591 | -35.2214
24.4293 | -35.5596 | -35.2271
24.4354 | -35.56 | -35.2329
24.4415 | -35.5604 | -35.2387
24.4476 | -35.5608 | -35.2449
24.4537 | -35.5612 | -35.2504
24.4598 | -35.5617 | -35.2564
24.4659 | -35.5621 | -35.2621
24.472 | -35.5625 | -35.2681
24.4781 | -35.5629 | -35.2741
24.4842 | -35.5632 | -35.2797
24.4903 | -35.5636 | -35.2856
24.4964 | -35.564 | -35.2913
24.5025 | -35.5644 | -35.2971
24.5086 | -35.5649 | -35.3027
24.5147 | -35.5653 | -35.3088
24.5208 | -35.5657 | -35.314
24.5269 | -35.5662 | -35.3205
24.533 | -35.5667 | -35.3267
24.5391 | -35.5672 | -35.3323
24.5452 | -35.5676 | -35.3381
24.5513 | -35.5681 | -35.3437
24.5574 | -35.5686 | -35.3501
24.5635 | -35.569 | -35.3563
24.5696 | -35.5694 | -35.3629
24.5758 | -35.5698 | -35.3687
24.5819 | -35.5702 | -35.375
24.588 | -35.5706 | -35.3805
24.5941 | -35.571 | -35.3865
24.6002 | -35.5714 | -35.392
24.6063 | -35.5718 | -35.3978
24.6124 | -35.5723 | -35.404
24.6185 | -35.5727 | -35.4098
24.6246 | -35.5732 | -35.4152
24.6307 | -35.5735 | -35.4215
24.6368 | -35.5739 | -35.4272
24.6429 | -35.5742 | -35.433
24.649 | -35.5746 | -35.4392
24.6551 | -35.575 | -35.4446
24.6612 | -35.5753 | -35.451
24.6673 | -35.5758 | -35.4568
24.6734 | -35.5762 | -35.4625
24.6795 | -35.5766 | -35.4685
24.6856 | -35.577 | -35.4744
24.6917 | -35.5774 | -35.4799
24.6978 | -35.5779 | -35.4858
24.7039 | -35.5783 | -35.4914
24.71 | -35.5786 | -35.4974
24.7162 | -35.5791 | -35.5035
24.7223 | -35.5795 | -35.5093
24.7284 | -35.58 | -35.5138
24.7345 | -35.5804 | -35.5196
24.7406 | -35.5808 | -35.5249
24.7467 | -35.5812 | -35.5312
24.7528 | -35.5816 | -35.5374
24.7589 | -35.582 | -35.5435
24.765 | -35.5824 | -35.5488
24.7711 | -35.5828 | -35.5546
24.7772 | -35.5831 | -35.5601
24.7833 | -35.5835 | -35.5661
24.7894 | -35.584 | -35.5717
24.7955 | -35.5844 | -35.5784
24.8016 | -35.5848 | -35.5836
24.8077 | -35.5852 | -35.5899
24.8138 | -35.5856 | -35.5956
24.8199 | -35.586 | -35.6013
24.826 | -35.5865 | -35.6075
24.8321 | -35.5869 | -35.6132
24.8382 | -35.5873 | -35.6193
24.8443 | -35.5877 | -35.6251
24.8504 | -35.5881 | -35.6306
24.8565 | -35.5885 | -35.6362
24.8627 | -35.5888 | -35.6424
24.8688 | -35.5892 | -35.6472
24.8749 | -35.5896 | -35.6531
24.881 | -35.5899 | -35.6587
24.8871 | -35.5903 | -35.6651
24.8932 | -35.5907 | -35.6712
24.8993 | -35.5912 | -35.6767
24.9054 | -35.5916 | -35.6829
24.9115 | -35.592 | -35.6882
24.9176 | -35.5924 | -35.6943
24.9237 | -35.5929 | -35.7
24.9298 | -35.5934 | -35.7055
24.9359 | -35.5938 | -35.7106
24.942 | -35.5941 | -35.7164
24.9481 | -35.5946 | -35.7221
24.9542 | -35.595 | -35.7281
24.9603 | -35.5954 | -35.7346
24.9664 | -35.5959 | -35.74
24.9725 | -35.5963 | -35.7458
24.9786 | -35.5966 | -35.7508
24.9847 | -35.5969 | -35.7566
24.9908 | -35.5974 | -35.7621
24.9969 | -35.5978 | -35.7681
25.0031 | -35.5982 | -35.7736
25.0092 | -35.5985 | -35.7794
25.0153 | -35.599 | -35.785
25.0214 | -35.5993 | -35.7911
25.0275 | -35.5998 | -35.7967
25.0336 | -35.6002 | -35.8023
25.0397 | -35.6006 | -35.8082
25.0458 | -35.6008 | -35.814
25.0519 | -35.6013 | -35.8194
25.058 | -35.6017 | -35.8253
25.0641 | -35.6021 | -35.8312
25.0702 | -35.6025 | -35.8377
25.0763 | -35.6029 | -35.8434
25.0824 | -35.6032 | -35.8487
25.0885 | -35.6036 | -35.8539
25.0946 | -35.604 | -35.8598
25.1007 | -35.6044 | -35.8659
25.1068 | -35.6049 | -35.8722
25.1129 | -35.6052 | -35.8785
25.119 | -35.6054 | -35.8841
25.1251 | -35.6057 | -35.8899
25.1312 | -35.6061 | -35.8958
25.1373 | -35.6065 | -35.902
25.1435 | -35.6068 | -35.9081
25.1496 | -35.6071 | -35.9133
25.1557 | -35.6074 | -35.9197
25.1618 | -35.6078 | -35.9255
25.1679 | -35.6082 | -35.9311
25.174 | -35.6085 | -35.9373
25.1801 | -35.6088 | -35.9429
25.1862 | -35.6093 | -35.9479
25.1923 | -35.6096 | -35.9533
25.1984 | -35.61 | -35.9593
25.2045 | -35.6104 | -35.9652
25.2106 | -35.6107 | -35.972
25.2167 | -35.6111 | -35.9779
25.2228 | -35.6115 | -35.9833
25.2289 | -35.612 | -35.9884
25.235 | -35.6124 | -35.9938
25.2411 | -35.6127 | -36
25.2472 | -35.6132 | -36.006
25.2533 | -35.6135 | -36.0123
25.2594 | -35.6139 | -36.0173
25.2655 | -35.6143 | -36.0234
25.2716 | -35.6146 | -36.0287
25.2777 | -35.6151 | -36.0351
25.2838 | -35.6155 | -36.0408
25.29 | -35.6158 | -36.046
25.2961 | -35.6161 | -36.0523
25.3022 | -35.6165 | -36.0577
25.3083 | -35.617 | -36.0633
25.3144 | -35.6174 | -36.0689
25.3205 | -35.6178 | -36.0751
25.3266 | -35.6182 | -36.0799
25.3327 | -35.6187 | -36.0856
25.3388 | -35.6191 | -36.0909
25.3449 | -35.6195 | -36.0966
25.351 | -35.6199 | -36.1029
25.3571 | -35.6204 | -36.109
25.3632 | -35.6207 | -36.1152
25.3693 | -35.6211 | -36.1211
25.3754 | -35.6215 | -36.1269
25.3815 | -35.622 | -36.1328
25.3876 | -35.6224 | -36.1378
25.3937 | -35.6228 | -36.1436
25.3998 | -35.6233 | -36.1495
25.4059 | -35.6237 | -36.1551
25.412 | -35.6241 | -36.1612
25.4181 | -35.6245 | -36.1671
25.4242 | -35.6249 | -36.1731
25.4304 | -35.6254 | -36.1783
25.4365 | -35.6258 | -36.1839
25.4426 | -35.6261 | -36.1892
25.4487 | -35.6265 | -36.1951
25.4548 | -35.6269 | -36.2008
25.4609 | -35.6272 | -36.2061
25.467 | -35.6276 | -36.2113
25.4731 | -35.628 | -36.2165
25.4792 | -35.6284 | -36.2222
25.4853 | -35.6287 | -36.2282
25.4914 | -35.6292 | -36.2341
25.4975 | -35.6295 | -36.24
25.5036 | -35.6299 | -36.2464
25.5097 | -35.6303 | -36.2519
25.5158 | -35.6307 | -36.257
25.5219 | -35.6311 | -36.2628
25.528 | -35.6314 | -36.2691
25.5341 | -35.6318 | -36.2745
25.5402 | -35.6322 | -36.2803
25.5463 | -35.6327 | -36.2857
25.5524 | -35.6331 | -36.291
25.5585 | -35.6334 | -36.2965
25.5646 | -35.6338 | -36.3024
25.5707 | -35.6343 | -36.3076
25.5769 | -35.6346 | -36.3124
25.583 | -35.635 | -36.3185
25.5891 | -35.6354 | -36.324
25.5952 | -35.6359 | -36.3299
25.6013 | -35.6362 | -36.3353
25.6074 | -35.6365 | -36.3404
25.6135 | -35.6368 | -36.3454
25.6196 | -35.6373 | -36.3509
25.6257 | -35.6376 | -36.3566
25.6318 | -35.638 | -36.3628
25.6379 | -35.6383 | -36.3686
25.644 | -35.6388 | -36.3741
25.6501 | -35.6392 | -36.3797
25.6562 | -35.6396 | -36.3854
25.6623 | -35.64 | -36.3905
25.6684 | -35.6404 | -36.3961
25.6745 | -35.6409 | -36.4025
25.6806 | -35.6412 | -36.4082
25.6867 | -35.6415 | -36.4137
25.6928 | -35.6419 | -36.4198
25.6989 | -35.6423 | -36.4257
25.705 | -35.6427 | -36.4307
25.7111 | -35.6431 | -36.4359
25.7173 | -35.6434 | -36.441
25.7234 | -35.6437 | -36.4469
25.7295 | -35.6441 | -36.4527
25.7356 | -35.6445 | -36.4587
25.7417 | -35.6449 | -36.4639
25.7478 | -35.6452 | -36.47
25.7539 | -35.6455 | -36.4749
25.76 | -35.6459 | -36.4811
25.7661 | -35.6462 | -36.4867
25.7722 | -35.6466 | -36.4921
25.7783 | -35.647 | -36.498
25.7844 | -35.6475 | -36.5032
25.7905 | -35.6479 | -36.5087
25.7966 | -35.6484 | -36.5141
25.8027 | -35.6488 | -36.5192
25.8088 | -35.6491 | -36.5251
25.8149 | -35.6495 | -36.53
25.821 | -35.6498 | -36.5356
25.8271 | -35.6502 | -36.541
25.8332 | -35.6506 | -36.5468
25.8393 | -35.651 | -36.5522
25.8454 | -35.6515 | -36.5576
25.8515 | -35.6519 | -36.5631
25.8576 | -35.6522 | -36.5684
25.8638 | -35.6527 | -36.5739
25.8699 | -35.6531 | -36.5794
25.876 | -35.6535 | -36.5843
25.8821 | -35.6539 | -36.5897
25.8882 | -35.6544 | -36.5946
25.8943 | -35.6548 | -36.6005
25.9004 | -35.6551 | -36.606
25.9065 | -35.6555 | -36.6117
25.9126 | -35.6559 | -36.6172
25.9187 | -35.6562 | -36.6228
25.9248 | -35.6566 | -36.6281
25.9309 | -35.657 | -36.6337
25.937 | -35.6574 | -36.6393
25.9431 | -35.6578 | -36.6448
25.9492 | -35.6581 | -36.6501
25.9553 | -35.6585 | -36.6553
25.9614 | -35.6588 | -36.6597
25.9675 | -35.6592 | -36.6655
25.9736 | -35.6596 | -36.6714
25.9797 | -35.6599 | -36.6769
25.9858 | -35.6602 | -36.6821
25.9919 | -35.6606 | -36.6876
25.998 | -35.6609 | -36.6931
26.0042 | -35.6613 | -36.6989
26.0103 | -35.6617 | -36.7043
26.0164 | -35.6621 | -36.7098
26.0225 | -35.6625 | -36.7148
26.0286 | -35.6629 | -36.7208
26.0347 | -35.6633 | -36.7253
26.0408 | -35.6637 | -36.7305
26.0469 | -35.6641 | -36.7355
26.053 | -35.6645 | -36.7407
26.0591 | -35.6649 | -36.7461
26.0652 | -35.6653 | -36.7516
26.0713 | -35.6656 | -36.7571
26.0774 | -35.666 | -36.7622
26.0835 | -35.6664 | -36.7674
26.0896 | -35.6668 | -36.7725
26.0957 | -35.6671 | -36.7778
26.1018 | -35.6675 | -36.7829
26.1079 | -35.6679 | -36.7888
26.114 | -35.6682 | -36.794
26.1201 | -35.6686 | -36.7995
26.1262 | -35.6689 | -36.8046
26.1323 | -35.6692 | -36.8096
26.1384 | -35.6698 | -36.8143
26.1445 | -35.6702 | -36.8201
26.1507 | -35.6707 | -36.8256
26.1568 | -35.671 | -36.8309
26.1629 | -35.6714 | -36.8362
26.169 | -35.6717 | -36.8418
26.1751 | -35.6721 | -36.847
26.1812 | -35.6724 | -36.8526
26.1873 | -35.6728 | -36.858
26.1934 | -35.6732 | -36.8631
26.1995 | -35.6737 | -36.8683
26.2056 | -35.6741 | -36.8743
26.2117 | -35.6744 | -36.8795
26.2178 | -35.6748 | -36.8849
26.2239 | -35.6751 | -36.8905
26.23 | -35.6755 | -36.8963
26.2361 | -35.6759 | -36.9024
26.2422 | -35.6763 | -36.9077
26.2483 | -35.6767 | -36.913
26.2544 | -35.677 | -36.9182
26.2605 | -35.6774 | -36.9237
26.2666 | -35.6777 | -36.9289
26.2727 | -35.6781 | -36.935
26.2788 | -35.6784 | -36.941
26.2849 | -35.6788 | -36.9467
26.2911 | -35.6793 | -36.9518
26.2972 | -35.6797 | -36.957
26.3033 | -35.6801 | -36.963
26.3094 | -35.6804 | -36.968
26.3155 | -35.6809 | -36.9726
26.3216 | -35.6812 | -36.9777
26.3277 | -35.6816 | -36.9831
26.3338 | -35.682 | -36.9881
26.3399 | -35.6823 | -36.9936
26.346 | -35.6827 | -36.9985
26.3521 | -35.683 | -37.0042
26.3582 | -35.6835 | -37.0096
26.3643 | -35.6839 | -37.0147
26.3704 | -35.6843 | -37.0199
26.3765 | -35.6847 | -37.0254
26.3826 | -35.685 | -37.0302
26.3887 | -35.6854 | -37.0358
26.3948 | -35.6858 | -37.0407
26.4009 | -35.6862 | -37.0459
26.407 | -35.6866 | -37.0512
26.4131 | -35.6869 | -37.0562
26.4192 | -35.6872 | -37.0615
26.4253 | -35.6876 | -37.0666
26.4314 | -35.688 | -37.0722
26.4376 | -35.6884 | -37.0781
26.4437 | -35.6887 | -37.0834
26.4498 | -35.6891 | -37.0891
26.4559 | -35.6894 | -37.0939
26.462 | -35.6897 | -37.0996
26.4681 | -35.69 | -37.105
26.4742 | -35.6904 | -37.11
26.4803 | -35.6908 | -37.1159
26.4864 | -35.6911 | -37.121
26.4925 | -35.6915 | -37.1263
26.4986 | -35.6919 | -37.1319
26.5047 | -35.6922 | -37.137
26.5108 | -35.6925 | -37.1416
26.5169 | -35.6929 | -37.1466
26.523 | -35.6933 | -37.15
26.5291 | -35.6937 | -37.1552
26.5352 | -35.6941 | -37.1605
26.5413 | -35.6945 | -37.1658
26.5474 | -35.6949 | -37.1714
26.5535 | -35.6952 | -37.1765
26.5596 | -35.6957 | -37.1817
26.5657 | -35.6961 | -37.1869
26.5718 | -35.6964 | -37.1928
26.578 | -35.6968 | -37.1981
26.5841 | -35.6972 | -37.203
26.5902 | -35.6976 | -37.2083
26.5963 | -35.6982 | -37.2135
26.6024 | -35.6986 | -37.2184
26.6085 | -35.699 | -37.2235
26.6146 | -35.6994 | -37.2286
26.6207 | -35.6998 | -37.2334
26.6268 | -35.7002 | -37.2386
26.6329 | -35.7006 | -37.2433
26.639 | -35.701 | -37.249
26.6451 | -35.7013 | -37.2541
26.6512 | -35.7017 | -37.2594
26.6573 | -35.7021 | -37.2645
26.6634 | -35.7024 | -37.2692
26.6695 | -35.7027 | -37.275
26.6756 | -35.703 | -37.2803
26.6817 | -35.7033 | -37.2864
26.6878 | -35.7038 | -37.2915
26.6939 | -35.7042 | -37.2973
26.7 | -35.7045 | -37.3018
26.7061 | -35.7049 | -37.3078
26.7122 | -35.7053 | -37.3136
26.7183 | -35.7056 | -37.3192
26.7245 | -35.706 | -37.324
26.7306 | -35.7065 | -37.3294
26.7367 | -35.7067 | -37.3343
26.7428 | -35.7071 | -37.3396
26.7489 | -35.7076 | -37.3448
26.755 | -35.708 | -37.3495
26.7611 | -35.7083 | -37.3551
26.7672 | -35.7087 | -37.3606
26.7733 | -35.7091 | -37.3653
26.7794 | -35.7096 | -37.3708
26.7855 | -35.7099 | -37.3761
26.7916 | -35.7103 | -37.3814
26.7977 | -35.7106 | -37.386
26.8038 | -35.7109 | -37.3904
26.8099 | -35.7113 | -37.3959
26.816 | -35.7115 | -37.4009
26.8221 | -35.7119 | -37.4055
26.8282 | -35.7123 | -37.4103
26.8343 | -35.7127 | -37.4152
26.8404 | -35.7131 | -37.4204
26.8465 | -35.7134 | -37.4259
26.8526 | -35.7139 | -37.4311
26.8587 | -35.7142 | -37.437
26.8649 | -35.7146 | -37.4416
26.871 | -35.715 | -37.4464
26.8771 | -35.7152 | -37.4512
26.8832 | -35.7156 | -37.4569
26.8893 | -35.716 | -37.4622
26.8954 | -35.7163 | -37.4675
26.9015 | -35.7166 | -37.4726
26.9076 | -35.717 | -37.4779
26.9137 | -35.7174 | -37.4831
26.9198 | -35.7178 | -37.4885
26.9259 | -35.7181 | -37.4941
26.932 | -35.7185 | -37.4996
26.9381 | -35.7188 | -37.5043
26.9442 | -35.7192 | -37.5087
26.9503 | -35.7197 | -37.5149
26.9564 | -35.72 | -37.5204
26.9625 | -35.7204 | -37.5256
26.9686 | -35.7208 | -37.5306
26.9747 | -35.7211 | -37.5355
26.9808 | -35.7215 | -37.5407
26.9869 | -35.7218 | -37.5458
26.993 | -35.7222 | -37.5504
26.9991 | -35.7225 | -37.5554
27.0052 | -35.7229 | -37.5612
27.0114 | -35.7232 | -37.5672
27.0175 | -35.7237 | -37.5725
27.0236 | -35.724 | -37.5771
27.0297 | -35.7243 | -37.5822
27.0358 | -35.7246 | -37.5874
27.0419 | -35.725 | -37.5933
27.048 | -35.7254 | -37.5981
27.0541 | -35.7258 | -37.6032
27.0602 | -35.7262 | -37.6079
27.0663 | -35.7265 | -37.6136
27.0724 | -35.727 | -37.6185
27.0785 | -35.7274 | -37.6242
27.0846 | -35.7276 | -37.6289
27.0907 | -35.728 | -37.6341
27.0968 | -35.7284 | -37.6389
27.1029 | -35.7287 | -37.6446
27.109 | -35.7291 | -37.6501
27.1151 | -35.7295 | -37.6558
27.1212 | -35.7298 | -37.6597
27.1273 | -35.7302 | -37.6649
27.1334 | -35.7307 | -37.6704
27.1395 | -35.731 | -37.6755
27.1456 | -35.7313 | -37.6802
27.1518 | -35.7316 | -37.6848
27.1579 | -35.7319 | -37.6889
27.164 | -35.7322 | -37.694
27.1701 | -35.7325 | -37.6993
27.1762 | -35.7329 | -37.7048
27.1823 | -35.7332 | -37.7096
27.1884 | -35.7336 | -37.7148
27.1945 | -35.7338 | -37.7202
27.2006 | -35.7343 | -37.7251
27.2067 | -35.7347 | -37.7297
27.2128 | -35.7351 | -37.7355
27.2189 | -35.7355 | -37.7405
27.225 | -35.7359 | -37.7462
27.2311 | -35.7363 | -37.7507
27.2372 | -35.7367 | -37.7552
27.2433 | -35.7371 | -37.7601
27.2494 | -35.7375 | -37.765
27.2555 | -35.7379 | -37.7703
27.2616 | -35.7382 | -37.7754
27.2677 | -35.7385 | -37.7802
27.2738 | -35.739 | -37.7858
27.2799 | -35.7393 | -37.7909
27.286 | -35.7398 | -37.7956
27.2921 | -35.7402 | -37.802
27.2983 | -35.7405 | -37.8069
27.3044 | -35.7408 | -37.8118
27.3105 | -35.741 | -37.8162
27.3166 | -35.7414 | -37.8214
27.3227 | -35.7418 | -37.8273
27.3288 | -35.7421 | -37.8331
27.3349 | -35.7425 | -37.8385
27.341 | -35.7429 | -37.8434
27.3471 | -35.7433 | -37.8489
27.3532 | -35.7437 | -37.8543
27.3593 | -35.7441 | -37.8598
27.3654 | -35.7445 | -37.8655
27.3715 | -35.7449 | -37.8702
27.3776 | -35.7453 | -37.8753
27.3837 | -35.7457 | -37.8799
27.3898 | -35.746 | -37.8844
27.3959 | -35.7463 | -37.8894
27.402 | -35.7467 | -37.8942
27.4081 | -35.7469 | -37.8995
27.4142 | -35.7473 | -37.9043
27.4203 | -35.7477 | -37.9096
27.4264 | -35.748 | -37.9143
27.4325 | -35.7484 | -37.9195
27.4387 | -35.7487 | -37.9245
27.4448 | -35.7489 | -37.9293
27.4509 | -35.7494 | -37.9349
27.457 | -35.7498 | -37.9396
27.4631 | -35.7501 | -37.9449
27.4692 | -35.7505 | -37.9501
27.4753 | -35.7509 | -37.956
27.4814 | -35.7512 | -37.9609
27.4875 | -35.7515 | -37.9667
27.4936 | -35.7518 | -37.9719
27.4997 | -35.7522 | -37.9768
27.5058 | -35.7524 | -37.9823
27.5119 | -35.7527 | -37.9875
27.518 | -35.7531 | -37.9927
27.5241 | -35.7535 | -37.9979
27.5302 | -35.7539 | -38.003
27.5363 | -35.7543 | -38.0075
27.5424 | -35.7546 | -38.0127
27.5485 | -35.7549 | -38.0178
27.5546 | -35.7553 | -38.0235
27.5607 | -35.7556 | -38.029
27.5668 | -35.756 | -38.0342
27.5729 | -35.7563 | -38.0393
27.5791 | -35.7567 | -38.0441
27.5852 | -35.757 | -38.0488
27.5913 | -35.7574 | -38.0537
27.5974 | -35.7579 | -38.059
27.6035 | -35.7582 | -38.0641
27.6096 | -35.7586 | -38.0694
27.6157 | -35.7589 | -38.0739
27.6218 | -35.7592 | -38.0789
27.6279 | -35.7596 | -38.0847
27.634 | -35.7598 | -38.0901
27.6401 | -35.7601 | -38.0953
27.6462 | -35.7605 | -38.1005
27.6523 | -35.7608 | -38.1057
27.6584 | -35.7612 | -38.1106
27.6645 | -35.7616 | -38.1159
27.6706 | -35.762 | -38.1211
27.6767 | -35.7623 | -38.1261
27.6828 | -35.7625 | -38.1312
27.6889 | -35.7628 | -38.1366
27.695 | -35.7631 | -38.1415
27.7011 | -35.7634 | -38.1464
27.7072 | -35.7638 | -38.1511
27.7133 | -35.7641 | -38.156
27.7194 | -35.7644 | -38.16
27.7256 | -35.7648 | -38.1649
27.7317 | -35.7652 | -38.1701
27.7378 | -35.7656 | -38.1756
27.7439 | -35.7659 | -38.1809
27.75 | -35.7663 | -38.1865
27.7561 | -35.7666 | -38.1916
27.7622 | -35.767 | -38.1963
27.7683 | -35.7672 | -38.2012
27.7744 | -35.7677 | -38.2061
27.7805 | -35.768 | -38.2109
27.7866 | -35.7684 | -38.2163
27.7927 | -35.7687 | -38.2219
27.7988 | -35.769 | -38.2269
27.8049 | -35.7694 | -38.2323
27.811 | -35.7698 | -38.2376
27.8171 | -35.77 | -38.2432
27.8232 | -35.7704 | -38.2491
27.8293 | -35.7709 | -38.2544
27.8354 | -35.7713 | -38.26
27.8415 | -35.7718 | -38.2645
27.8476 | -35.772 | -38.2703
27.8537 | -35.7724 | -38.2756
27.8598 | -35.7728 | -38.2806
27.866 | -35.7731 | -38.2853
27.8721 | -35.7735 | -38.2905
27.8782 | -35.7738 | -38.2959
27.8843 | -35.7741 | -38.3017
27.8904 | -35.7745 | -38.3067
27.8965 | -35.7748 | -38.3115
27.9026 | -35.7752 | -38.3164
27.9087 | -35.7755 | -38.3214
27.9148 | -35.7758 | -38.3263
27.9209 | -35.7762 | -38.3313
27.927 | -35.7765 | -38.3363
27.9331 | -35.7768 | -38.3414
27.9392 | -35.7771 | -38.3461
27.9453 | -35.7775 | -38.3516
27.9514 | -35.7779 | -38.357
27.9575 | -35.7782 | -38.3618
27.9636 | -35.7785 | -38.3669
27.9697 | -35.7788 | -38.3719
27.9758 | -35.7791 | -38.3768
27.9819 | -35.7794 | -38.3822
27.988 | -35.7798 | -38.3867
27.9941 | -35.7801 | -38.3918
28.0002 | -35.7805 | -38.397
28.0063 | -35.7808 | -38.4022
28.0125 | -35.7811 | -38.4076
28.0186 | -35.7815 | -38.4122
28.0247 | -35.7818 | -38.4173
28.0308 | -35.7822 | -38.4227
28.0369 | -35.7826 | -38.4274
28.043 | -35.7829 | -38.4327
28.0491 | -35.7833 | -38.4378
28.0552 | -35.7835 | -38.4434
28.0613 | -35.784 | -38.4489
28.0674 | -35.7842 | -38.4541
28.0735 | -35.7847 | -38.4586

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.257 | -42.591
21.4908 | -35.4877 | -35.5222
16.4948 | -34.8613 | -28.025
12.6602 | -34.4929 | -21.6809
9.7171 | -34.2805 | -16.7051
7.4582 | -34.129 | -12.6853
5.7244 | -34.0215 | -9.427
4.3936 | -33.9598 | -6.8047
3.3722 | -33.944 | -4.7367
2.5883 | -33.938 | -3.1692
1.9866 | -33.9364 | -1.9176
1.5248 | -33.936 | -0.91402
1.1703 | -33.945 | -0.26164
1 | -33.9403 | 0.062051
0.89822 | -33.9444 | 0.33182
0.68941 | -33.9256 | 0.91519
0.52919 | -33.9264 | 1.3665
0.40612 | -33.9261 | 1.9753
0.31171 | -33.9193 | 2.4276
0.23928 | -33.9173 | 3.1507
0.18358 | -33.9284 | 3.5627
0.14091 | -33.943 | 4.4412
0.10819 | -33.9675 | 5.595
0.083017 | -33.9903 | 7.0706
0.063705 | -34.0316 | 8.9381
0.048923 | -34.1085 | 11.4281
0.037479 | -34.2177 | 14.6141
0.028801 | -34.3925 | 18.6243
0.022078 | -34.6903 | 23.9413
0.017023 | -35.1435 | 30.345
0.013018 | -35.9443 | 38.0207
0.010014 | -37.0527 | 45.9269

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003658 | -7.821
0.007316 | -24.7308
0.010974 | -34.1669
0.014633 | -40.5891
0.018291 | -43.9501
0.021949 | -45.8182
0.025607 | -47.369
0.029265 | -49.0419
0.032923 | -49.6026
0.036581 | -50.4102
0.04024 | -51.79
0.043898 | -53.6632
0.047556 | -54.9583
0.051214 | -56.0177
0.054872 | -56.4904
0.05853 | -58.8872
0.062188 | -59.6774
0.065847 | -59.507
0.069505 | -60.0657
0.073163 | -60.635
0.076821 | -60.161
0.080479 | -60.667
0.084137 | -61.1688
0.087796 | -61.5335
0.091454 | -61.7849
0.095112 | -62.7742
0.09877 | -63.5965
0.10243 | -64.0336
0.10609 | -63.3921
0.10974 | -65.0604
0.1134 | -66.5772
0.11706 | -66.5217
0.12072 | -66.6351
0.12438 | -67.6234
0.12804 | -67.1484
0.13169 | -67.0465
0.13535 | -67.1353
0.13901 | -68.1232
0.14267 | -68.6331
0.14633 | -68.2502
0.14998 | -67.8528
0.15364 | -68.1736
0.1573 | -68.4042
0.16096 | -69.6445
0.16462 | -69.0004
0.16828 | -69.8012
0.17193 | -70.4251
0.17559 | -71.3039
0.17925 | -71.7469
0.18291 | -70.2585
0.18657 | -70.1877
0.19022 | -70.3881
0.19388 | -71.4938
0.19754 | -72.6315
0.2012 | -72.4213
0.20486 | -71.681
0.20851 | -71.4891
0.21217 | -71.7189
0.21583 | -72.2599
0.21949 | -72.3901
0.22315 | -72.106
0.22681 | -73.3974
0.23046 | -74.3909
0.23412 | -72.5797
0.23778 | -71.556
0.24144 | -72.4578
0.2451 | -72.5001
0.24875 | -73.4017
0.25241 | -73.5589
0.25607 | -72.6975
0.25973 | -73.734
0.26339 | -75.1219
0.26704 | -75.3031
0.2707 | -74.8712
0.27436 | -75.0067
0.27802 | -74.9683
0.28168 | -75.7489
0.28534 | -75.6766
0.28899 | -75.2155
0.29265 | -75.8563
0.29631 | -76.3599
0.29997 | -75.8819
0.30363 | -74.5716
0.30728 | -75.9564
0.31094 | -76.2389
0.3146 | -77.0046
0.31826 | -76.8465
0.32192 | -76.579
0.32558 | -75.0795
0.32923 | -75.0781
0.33289 | -75.4452
0.33655 | -76.9977
0.34021 | -77.5933
0.34387 | -77.3966
0.34752 | -77.6237
0.35118 | -77.9694
0.35484 | -78.0505
0.3585 | -77.223
0.36216 | -76.7713
0.36582 | -76.8994
0.36947 | -76.9367
0.37313 | -77.2246
0.37679 | -77.7739
0.38045 | -78.9194
0.3841 | -79.1874
0.38776 | -78.9173
0.39142 | -78.3089
0.39508 | -78.2713
0.39874 | -78.1564
0.4024 | -79.3875
0.40605 | -79.3222
0.40971 | -79.3943
0.41337 | -79.083
0.41703 | -79.2529
0.42069 | -79.3273
0.42435 | -78.6954
0.428 | -78.9552
0.43166 | -78.3727
0.43532 | -78.3924
0.43898 | -78.5067
0.44264 | -79.0546
0.44629 | -78.1651
0.44995 | -77.77
0.45361 | -79.4545
0.45727 | -80.5691
0.46093 | -79.6474
0.46459 | -79.6909
0.46824 | -79.2983
0.4719 | -79.7767
0.47556 | -81.0935
0.47922 | -80.6107
0.48287 | -79.9885
0.48653 | -80.2268
0.49019 | -80.3334
0.49385 | -80.4901
0.49751 | -80.3391
0.50117 | -80.9127
0.50482 | -80.4881
0.50848 | -80.6935
0.51214 | -81.1158
0.5158 | -81.3671
0.51946 | -81.3494
0.52311 | -82.3346
0.52677 | -81.8333
0.53043 | -81.3914
0.53409 | -80.8988
0.53775 | -80.8376
0.54141 | -81.1103
0.54506 | -81.3573
0.54872 | -81.0534
0.55238 | -80.6747
0.55604 | -80.7036
0.5597 | -81.355
0.56336 | -82.1973
0.56701 | -82.6497
0.57067 | -82.9886
0.57433 | -82.9208
0.57799 | -82.1799
0.58164 | -81.3561
0.5853 | -81.1882
0.58896 | -81.5038
0.59262 | -81.5989
0.59628 | -81.6619
0.59994 | -81.9859
0.60359 | -82.0408
0.60725 | -81.847
0.61091 | -82.1486
0.61457 | -82.6993
0.61823 | -82.7728
0.62189 | -82.5054
0.62554 | -82.8139
0.6292 | -83.7883
0.63286 | -84.2629
0.63652 | -83.6441
0.64018 | -82.8713
0.64383 | -82.6041
0.64749 | -82.9974
0.65115 | -83.3144
0.65481 | -83.0964
0.65847 | -82.9001
0.66212 | -82.8826
0.66578 | -82.9219
0.66944 | -83.1241
0.6731 | -83.3178
0.67676 | -83.2682
0.68041 | -83.2778
0.68407 | -83.4488
0.68773 | -83.5904
0.69139 | -83.5338
0.69505 | -83.5288
0.69871 | -83.1902
0.70236 | -82.7777
0.70602 | -82.881
0.70968 | -83.4893
0.71334 | -83.6909
0.717 | -83.5384
0.72066 | -83.9483
0.72431 | -84.3893
0.72797 | -84.3513
0.73163 | -83.9197
0.73529 | -83.7257
0.73895 | -84.0784
0.7426 | -84.5902
0.74626 | -84.8659
0.74992 | -85.3993
0.75358 | -85.7713
0.75724 | -85.3197
0.76089 | -84.5487
0.76455 | -84.2753
0.76821 | -84.4912
0.77187 | -84.583
0.77553 | -84.3746
0.77919 | -84.554
0.78284 | -85.093
0.7865 | -85.5965
0.79016 | -85.6699
0.79382 | -85.4509
0.79748 | -85.4155
0.80113 | -85.3775
0.80479 | -85.2467
0.80845 | -85.5802
0.81211 | -86.1178
0.81577 | -86.3004
0.81942 | -86.1968
0.82308 | -85.8275
0.82674 | -85.3734
0.8304 | -85.3114
0.83406 | -85.9078
0.83771 | -86.5426
0.84137 | -86.5296
0.84503 | -85.4747
0.84869 | -84.7599
0.85235 | -85.0166
0.85601 | -85.6068
0.85966 | -85.9869
0.86332 | -85.914
0.86698 | -85.837
0.87064 | -85.894
0.8743 | -85.8794
0.87796 | -85.648
0.88161 | -85.5645
0.88527 | -85.8208
0.88893 | -86.306
0.89259 | -86.2962
0.89625 | -86.0458
0.8999 | -86.2187
0.90356 | -86.0746
0.90722 | -85.615
0.91088 | -85.6283
0.91454 | -85.9735
0.91819 | -86.2276
0.92185 | -86.1865
0.92551 | -86.1974
0.92917 | -86.5129
0.93283 | -86.7698
0.93649 | -86.5202
0.94014 | -86.2419
0.9438 | -86.4467
0.94746 | -87.3491
0.95112 | -88.168
0.95478 | -87.8721
0.95843 | -87.55
0.96209 | -87.3883
0.96575 | -86.9749
0.96941 | -86.8236
0.97307 | -86.7983
0.97672 | -86.4154
0.98038 | -86.1378
0.98404 | -86.3684
0.9877 | -86.7316
0.99136 | -86.8399
0.99502 | -86.7112
0.99867 | -86.7408
1.0023 | -87.0464
1.006 | -87.3348
1.0096 | -87.1846
1.0133 | -87.078
1.017 | -87.1895
1.0206 | -87.1538
1.0243 | -87.3283
1.0279 | -87.9252
1.0316 | -88.6601
1.0353 | -89.0946
1.0389 | -89.0093
1.0426 | -88.6012
1.0462 | -88.0195
1.0499 | -87.5759
1.0535 | -87.6353
1.0572 | -87.7429
1.0609 | -87.8871
1.0645 | -87.8299
1.0682 | -87.6165
1.0718 | -87.5823
1.0755 | -87.8249
1.0792 | -88.1861
1.0828 | -88.5007
1.0865 | -88.7108
1.0901 | -88.7591
1.0938 | -88.6034
1.0974 | -88.4973
1.1011 | -88.621
1.1048 | -88.6768
1.1084 | -88.4838
1.1121 | -88.1991
1.1157 | -87.8819
1.1194 | -87.652
1.1231 | -87.7155
1.1267 | -88.129
1.1304 | -88.7362
1.134 | -89.0911
1.1377 | -88.8913
1.1413 | -88.4586
1.145 | -88.2918
1.1487 | -88.5205
1.1523 | -88.808
1.156 | -88.8402
1.1596 | -88.7641
1.1633 | -88.7315
1.1669 | -88.6968
1.1706 | -88.6242
1.1743 | -88.6183
1.1779 | -88.8813
1.1816 | -89.3559
1.1852 | -89.6807
1.1889 | -89.7146
1.1926 | -89.5793
1.1962 | -89.3961
1.1999 | -89.3582
1.2035 | -89.4113
1.2072 | -89.29
1.2108 | -88.9992
1.2145 | -88.8386
1.2182 | -88.9883
1.2218 | -89.2485
1.2255 | -89.2977
1.2291 | -89.228
1.2328 | -89.2993
1.2365 | -89.4632
1.2401 | -89.4373
1.2438 | -89.2335
1.2474 | -89.231
1.2511 | -89.5417
1.2547 | -89.8538
1.2584 | -90.0197
1.2621 | -90.2207
1.2657 | -90.418
1.2694 | -90.2626
1.273 | -89.7889
1.2767 | -89.4796
1.2804 | -89.5624
1.284 | -89.9554
1.2877 | -90.4789
1.2913 | -90.8391
1.295 | -90.8434
1.2986 | -90.678
1.3023 | -90.5117
1.306 | -90.291
1.3096 | -90.0991
1.3133 | -90.0666
1.3169 | -89.9492
1.3206 | -89.4878
1.3242 | -88.9439
1.3279 | -88.7823
1.3316 | -89.1979
1.3352 | -89.9362
1.3389 | -90.4141
1.3425 | -90.5466
1.3462 | -90.6377
1.3499 | -90.6646
1.3535 | -90.5358
1.3572 | -90.3142
1.3608 | -90.0292
1.3645 | -89.7532
1.3681 | -89.5767
1.3718 | -89.4978
1.3755 | -89.4538
1.3791 | -89.5332
1.3828 | -89.8663
1.3864 | -90.2781
1.3901 | -90.4446
1.3938 | -90.4464
1.3974 | -90.6048
1.4011 | -90.8382
1.4047 | -90.7773
1.4084 | -90.4984
1.412 | -90.3189
1.4157 | -90.3002
1.4194 | -90.3999
1.423 | -90.611
1.4267 | -90.8722
1.4303 | -91.0342
1.434 | -91.0268
1.4377 | -90.9008
1.4413 | -90.7384
1.445 | -90.7349
1.4486 | -90.9737
1.4523 | -91.3664
1.4559 | -91.7922
1.4596 | -91.9386
1.4633 | -91.5866
1.4669 | -91.0212
1.4706 | -90.6751
1.4742 | -90.7541
1.4779 | -91.0623
1.4815 | -91.2692
1.4852 | -91.3964
1.4889 | -91.5999
1.4925 | -91.7475
1.4962 | -91.5957
1.4998 | -91.2631
1.5035 | -91.0961
1.5072 | -91.0942
1.5108 | -91.0654
1.5145 | -91.0662
1.5181 | -91.0466
1.5218 | -90.7905
1.5254 | -90.4666
1.5291 | -90.3495
1.5328 | -90.4557
1.5364 | -90.7192
1.5401 | -91.0518
1.5437 | -91.3158
1.5474 | -91.5387
1.5511 | -91.8415
1.5547 | -92.1449
1.5584 | -92.258
1.562 | -92.1336
1.5657 | -91.9542
1.5693 | -91.8532
1.573 | -91.7545
1.5767 | -91.673
1.5803 | -91.7326
1.584 | -91.72
1.5876 | -91.6331
1.5913 | -91.5214
1.595 | -91.4746
1.5986 | -91.5331
1.6023 | -91.6691
1.6059 | -91.7661
1.6096 | -91.736
1.6132 | -91.6381
1.6169 | -91.588
1.6206 | -91.6332
1.6242 | -91.7238
1.6279 | -91.8055
1.6315 | -91.9009
1.6352 | -92.0746
1.6388 | -92.3093
1.6425 | -92.4511
1.6462 | -92.3861
1.6498 | -92.1971
1.6535 | -92.0146
1.6571 | -91.868
1.6608 | -91.7271
1.6645 | -91.581
1.6681 | -91.4863
1.6718 | -91.5036
1.6754 | -91.6059
1.6791 | -91.7064
1.6827 | -91.7211
1.6864 | -91.7155
1.6901 | -91.822
1.6937 | -92.0811
1.6974 | -92.3852
1.701 | -92.5423
1.7047 | -92.4458
1.7084 | -92.196
1.712 | -91.9991
1.7157 | -91.9963
1.7193 | -92.1902
1.723 | -92.4768
1.7266 | -92.7333
1.7303 | -92.881
1.734 | -92.8865
1.7376 | -92.783
1.7413 | -92.6311
1.7449 | -92.4675
1.7486 | -92.2894
1.7523 | -92.0888
1.7559 | -91.8933
1.7596 | -91.7349
1.7632 | -91.6773
1.7669 | -91.7846
1.7705 | -92.033
1.7742 | -92.3071
1.7779 | -92.535
1.7815 | -92.7261
1.7852 | -92.9234
1.7888 | -93.0528
1.7925 | -93.0146
1.7961 | -92.7842
1.7998 | -92.4821
1.8035 | -92.2827
1.8071 | -92.2969
1.8108 | -92.4949
1.8144 | -92.7511
1.8181 | -92.9016
1.8218 | -92.8933
1.8254 | -92.7629
1.8291 | -92.5913
1.8327 | -92.4264
1.8364 | -92.2889
1.84 | -92.2086
1.8437 | -92.2031
1.8474 | -92.2708
1.851 | -92.4143
1.8547 | -92.6052
1.8583 | -92.7717
1.862 | -92.886
1.8657 | -92.9821
1.8693 | -93.1145
1.873 | -93.2181
1.8766 | -93.1598
1.8803 | -92.9484
1.8839 | -92.733
1.8876 | -92.6069
1.8913 | -92.5353
1.8949 | -92.4651
1.8986 | -92.4076
1.9022 | -92.399
1.9059 | -92.4639
1.9096 | -92.5771
1.9132 | -92.6883
1.9169 | -92.7972
1.9205 | -92.9279
1.9242 | -93.0429
1.9278 | -93.0841
1.9315 | -93.0576
1.9352 | -93.05
1.9388 | -93.1141
1.9425 | -93.2002
1.9461 | -93.2376
1.9498 | -93.2221
1.9534 | -93.2061
1.9571 | -93.2427
1.9608 | -93.3706
1.9644 | -93.5618
1.9681 | -93.6746
1.9717 | -93.6079
1.9754 | -93.4093
1.9791 | -93.2316
1.9827 | -93.1683
1.9864 | -93.2339
1.99 | -93.4046
1.9937 | -93.5972
1.9973 | -93.7382
2.001 | -93.8162
2.0047 | -93.8345
2.0083 | -93.7891
2.012 | -93.7064
2.0156 | -93.6688
2.0193 | -93.7042
2.023 | -93.7656
2.0266 | -93.7982
2.0303 | -93.7702
2.0339 | -93.659
2.0376 | -93.4791
2.0412 | -93.2968
2.0449 | -93.2354
2.0486 | -93.3665
2.0522 | -93.6034
2.0559 | -93.7961
2.0595 | -93.884
2.0632 | -93.895
2.0669 | -93.8513
2.0705 | -93.7696
2.0742 | -93.6789
2.0778 | -93.6089
2.0815 | -93.6029
2.0851 | -93.6883
2.0888 | -93.8357
2.0925 | -94.0084
2.0961 | -94.1582
2.0998 | -94.2453
2.1034 | -94.2447
2.1071 | -94.2046
2.1107 | -94.2005
2.1144 | -94.1829
2.1181 | -94.1363
2.1217 | -94.0391
2.1254 | -93.9106
2.129 | -93.8006
2.1327 | -93.7488
2.1364 | -93.7553
2.14 | -93.7949
2.1437 | -93.8206
2.1473 | -93.8355
2.151 | -93.8232
2.1546 | -93.7671
2.1583 | -93.6745
2.162 | -93.572
2.1656 | -93.4969
2.1693 | -93.4711
2.1729 | -93.5149
2.1766 | -93.6353
2.1803 | -93.8306
2.1839 | -94.0858
2.1876 | -94.3303
2.1912 | -94.527
2.1949 | -94.6703
2.1985 | -94.7446
2.2022 | -94.7591
2.2059 | -94.724
2.2095 | -94.6488
2.2132 | -94.5176
2.2168 | -94.3313
2.2205 | -94.1261
2.2242 | -93.96
2.2278 | -93.8983
2.2315 | -93.953
2.2351 | -94.1003
2.2388 | -94.2907
2.2424 | -94.4341
2.2461 | -94.4806
2.2498 | -94.4317
2.2534 | -94.3404
2.2571 | -94.2455
2.2607 | -94.1642
2.2644 | -94.0938
2.2681 | -94.0308
2.2717 | -93.9941
2.2754 | -93.969
2.279 | -93.9466
2.2827 | -93.9474
2.2863 | -93.9747
2.29 | -94.0237
2.2937 | -94.0736
2.2973 | -94.0937
2.301 | -94.1035
2.3046 | -94.158
2.3083 | -94.2803
2.3119 | -94.4587
2.3156 | -94.6596
2.3193 | -94.8263
2.3229 | -94.9091
2.3266 | -94.9143
2.3302 | -94.8582
2.3339 | -94.75
2.3376 | -94.6438
2.3412 | -94.5774
2.3449 | -94.5559
2.3485 | -94.5305
2.3522 | -94.444
2.3558 | -94.2912
2.3595 | -94.13
2.3632 | -93.9975
2.3668 | -93.8866
2.3705 | -93.8118
2.3741 | -93.82
2.3778 | -93.9515
2.3815 | -94.1982
2.3851 | -94.5095
2.3888 | -94.81
2.3924 | -95.0436
2.3961 | -95.1837
2.3997 | -95.1882
2.4034 | -95.0893
2.4071 | -94.9607
2.4107 | -94.859
2.4144 | -94.8032
2.418 | -94.7727
2.4217 | -94.7462
2.4254 | -94.7345
2.429 | -94.7962
2.4327 | -94.9753
2.4363 | -95.2814
2.44 | -95.6549
2.4436 | -95.9509
2.4473 | -96.0809
2.451 | -96.0598
2.4546 | -95.9288
2.4583 | -95.7106
2.4619 | -95.4422
2.4656 | -95.155
2.4692 | -94.9103
2.4729 | -94.7582
2.4766 | -94.732
2.4802 | -94.8345
2.4839 | -95.0249
2.4875 | -95.2211
2.4912 | -95.3568
2.4949 | -95.4396
2.4985 | -95.5218
2.5022 | -95.6207
2.5058 | -95.7432
2.5095 | -95.8409
2.5131 | -95.886
2.5168 | -95.8548
2.5205 | -95.7411
2.5241 | -95.5743
2.5278 | -95.3803
2.5314 | -95.2087
2.5351 | -95.1199
2.5388 | -95.1212
2.5424 | -95.1943
2.5461 | -95.2752
2.5497 | -95.3413
2.5534 | -95.3901
2.557 | -95.4239
2.5607 | -95.4128
2.5644 | -95.3283
2.568 | -95.182
2.5717 | -95.0126
2.5753 | -94.8763
2.579 | -94.8126
2.5827 | -94.8221
2.5863 | -94.8719
2.59 | -94.9188
2.5936 | -94.9512
2.5973 | -94.9944
2.6009 | -95.0817
2.6046 | -95.2109
2.6083 | -95.3526
2.6119 | -95.46
2.6156 | -95.5064
2.6192 | -95.5148
2.6229 | -95.5224
2.6265 | -95.5638
2.6302 | -95.6232
2.6339 | -95.6692
2.6375 | -95.6189
2.6412 | -95.5875
2.6448 | -95.5517
2.6485 | -95.5351
2.6522 | -95.551
2.6558 | -95.5854
2.6595 | -95.614
2.6631 | -95.6258
2.6668 | -95.6105
2.6704 | -95.5818
2.6741 | -95.5648
2.6778 | -95.5586
2.6814 | -95.5663
2.6851 | -95.576
2.6887 | -95.5926
2.6924 | -95.612
2.6961 | -95.6264
2.6997 | -95.6433
2.7034 | -95.6658
2.707 | -95.6886
2.7107 | -95.693
2.7143 | -95.6754
2.718 | -95.6437
2.7217 | -95.6001
2.7253 | -95.5666
2.729 | -95.5549
2.7326 | -95.5656
2.7363 | -95.5973
2.74 | -95.6624
2.7436 | -95.7538
2.7473 | -95.8487
2.7509 | -95.9277
2.7546 | -95.9703
2.7582 | -95.9753
2.7619 | -95.9573
2.7656 | -95.9422
2.7692 | -95.9649
2.7729 | -96.0387
2.7765 | -96.1524
2.7802 | -96.2853
2.7838 | -96.4034
2.7875 | -96.4751
2.7912 | -96.4897
2.7948 | -96.4463
2.7985 | -96.3797
2.8021 | -96.3339
2.8058 | -96.33
2.8095 | -96.3718
2.8131 | -96.4229
2.8168 | -96.4661
2.8204 | -96.4905
2.8241 | -96.5057
2.8277 | -96.5256
2.8314 | -96.5537
2.8351 | -96.5747
2.8387 | -96.5918
2.8424 | -96.5965
2.846 | -96.5874
2.8497 | -96.5715
2.8534 | -96.5605
2.857 | -96.5632
2.8607 | -96.5906
2.8643 | -96.6177
2.868 | -96.6418
2.8716 | -96.6619
2.8753 | -96.6634
2.879 | -96.6349
2.8826 | -96.5453
2.8863 | -96.3618
2.8899 | -96.1095
2.8936 | -95.8395
2.8973 | -95.6034
2.9009 | -95.447
2.9046 | -95.3875
2.9082 | -95.4318
2.9119 | -95.551
2.9155 | -95.6816
2.9192 | -95.7764
2.9229 | -95.8167
2.9265 | -95.8182
2.9302 | -95.827
2.9338 | -95.8854
2.9375 | -96.0097
2.9411 | -96.2044
2.9448 | -96.4438
2.9485 | -96.6913
2.9521 | -96.9035
2.9558 | -97.0119
2.9594 | -96.9926
2.9631 | -96.8591
2.9668 | -96.6309
2.9704 | -96.3736
2.9741 | -96.1565
2.9777 | -96.0275
2.9814 | -96.0114
2.985 | -96.0939
2.9887 | -96.227
2.9924 | -96.3653
2.996 | -96.4533
2.9997 | -96.4826
3.0033 | -96.4877
3.007 | -96.5008
3.0107 | -96.5229
3.0143 | -96.5673
3.018 | -96.6231
3.0216 | -96.6607
3.0253 | -96.6886
3.0289 | -96.7057
3.0326 | -96.726
3.0363 | -96.736
3.0399 | -96.711
3.0436 | -96.6699
3.0472 | -96.6205
3.0509 | -96.5782
3.0546 | -96.5516
3.0582 | -96.5552
3.0619 | -96.5846
3.0655 | -96.6368
3.0692 | -96.7104
3.0728 | -96.7633
3.0765 | -96.7698
3.0802 | -96.7186
3.0838 | -96.6142
3.0875 | -96.4688
3.0911 | -96.2999
3.0948 | -96.1684
3.0985 | -96.1256
3.1021 | -96.1869
3.1058 | -96.3442
3.1094 | -96.5485
3.1131 | -96.7369
3.1167 | -96.8588
3.1204 | -96.8923
3.1241 | -96.87
3.1277 | -96.8426
3.1314 | -96.8689
3.135 | -96.9701
3.1387 | -97.1179
3.1423 | -97.2438
3.146 | -97.295
3.1497 | -97.242
3.1533 | -97.0873
3.157 | -96.8768
3.1606 | -96.6617
3.1643 | -96.6067
3.168 | -96.5492
3.1716 | -96.5534
3.1753 | -96.5989
3.1789 | -96.6485
3.1826 | -96.6596
3.1862 | -96.6051
3.1899 | -96.5011
3.1936 | -96.3741
3.1972 | -96.2379
3.2009 | -96.1321
3.2045 | -96.074
3.2082 | -96.0645
3.2119 | -96.1062
3.2155 | -96.1816
3.2192 | -96.2902
3.2228 | -96.4294
3.2265 | -96.5925
3.2301 | -96.7709
3.2338 | -96.9494
3.2375 | -97.1095
3.2411 | -97.2334
3.2448 | -97.3106
3.2484 | -97.3655
3.2521 | -97.4157
3.2557 | -97.4713
3.2594 | -97.5142
3.2631 | -97.542
3.2667 | -97.5507
3.2704 | -97.5515
3.274 | -97.5501
3.2777 | -97.5369
3.2814 | -97.5271
3.285 | -97.504
3.2887 | -97.4618
3.2923 | -97.4049
3.296 | -97.3567
3.2996 | -97.335
3.3033 | -97.3584
3.307 | -97.4296
3.3106 | -97.5199
3.3143 | -97.6182
3.3179 | -97.6929
3.3216 | -97.7115
3.3253 | -97.6711
3.3289 | -97.5927
3.3326 | -97.5083
3.3362 | -97.4155
3.3399 | -97.3443
3.3435 | -97.3023
3.3472 | -97.2873
3.3509 | -97.3069
3.3545 | -97.3633
3.3582 | -97.4525
3.3618 | -97.5728
3.3655 | -97.7187
3.3692 | -97.869
3.3728 | -98.0027
3.3765 | -98.11
3.3801 | -98.1662
3.3838 | -98.1523
3.3874 | -98.0758
3.3911 | -97.9496
3.3948 | -97.7985
3.3984 | -97.6605
3.4021 | -97.5702
3.4057 | -97.5307
3.4094 | -97.5438
3.4131 | -97.5998
3.4167 | -97.6859
3.4204 | -97.7634
3.424 | -97.8238
3.4277 | -97.8405
3.4313 | -97.8159
3.435 | -97.7757
3.4387 | -97.7314
3.4423 | -97.7
3.446 | -97.7114
3.4496 | -97.7665
3.4533 | -97.837
3.4569 | -97.9034
3.4606 | -97.9512
3.4643 | -97.9462
3.4679 | -97.8582
3.4716 | -97.696
3.4752 | -97.4978
3.4789 | -97.2861
3.4826 | -97.1094
3.4862 | -96.9951
3.4899 | -96.9594
3.4935 | -97.0123
3.4972 | -97.1428
3.5008 | -97.3213
3.5045 | -97.5236
3.5082 | -97.734
3.5118 | -97.9289
3.5155 | -98.0951
3.5191 | -98.2276
3.5228 | -98.3185
3.5265 | -98.368
3.5301 | -98.3838
3.5338 | -98.3705
3.5374 | -98.3396
3.5411 | -98.2946
3.5447 | -98.2328
3.5484 | -98.1488
3.5521 | -98.0503
3.5557 | -97.9386
3.5594 | -97.8172
3.563 | -97.7083
3.5667 | -97.6325
3.5704 | -97.6033
3.574 | -97.6349
3.5777 | -97.7172
3.5813 | -97.8295
3.585 | -97.9779
3.5886 | -98.1469
3.5923 | -98.3045
3.596 | -98.4421
3.5996 | -98.553
3.6033 | -98.6228
3.6069 | -98.647
3.6106 | -98.6373
3.6142 | -98.5902
3.6179 | -98.4982
3.6216 | -98.3735
3.6252 | -98.2414
3.6289 | -98.116
3.6325 | -98.0151
3.6362 | -97.9373
3.6399 | -97.881
3.6435 | -97.8589
3.6472 | -97.8723
3.6508 | -97.9095
3.6545 | -97.9747
3.6581 | -98.0548
3.6618 | -98.1262
3.6655 | -98.1856
3.6691 | -98.2291
3.6728 | -98.2577
3.6764 | -98.2881
3.6801 | -98.3281
3.6838 | -98.3823
3.6874 | -98.4548
3.6911 | -98.5577
3.6947 | -98.6417
3.6984 | -98.7265
3.702 | -98.8127
3.7057 | -98.8784
3.7094 | -98.9142
3.713 | -98.9136
3.7167 | -98.8657
3.7203 | -98.7869
3.724 | -98.7028
3.7277 | -98.6292
3.7313 | -98.5777
3.735 | -98.5545
3.7386 | -98.5634
3.7423 | -98.5998
3.7459 | -98.6401
3.7496 | -98.6835
3.7533 | -98.7244
3.7569 | -98.7519
3.7606 | -98.7525
3.7642 | -98.7088
3.7679 | -98.6279
3.7715 | -98.5196
3.7752 | -98.3835
3.7789 | -98.2502
3.7825 | -98.1153
3.7862 | -97.9973
3.7898 | -97.9257
3.7935 | -97.9099
3.7972 | -97.9452
3.8008 | -98.0309
3.8045 | -98.1437
3.8081 | -98.2721
3.8118 | -98.397
3.8154 | -98.4898
3.8191 | -98.5504
3.8228 | -98.5798
3.8264 | -98.5894
3.8301 | -98.5855
3.8337 | -98.5723
3.8374 | -98.5629
3.8411 | -98.562
3.8447 | -98.5737
3.8484 | -98.5882
3.852 | -98.5894
3.8557 | -98.5619
3.8593 | -98.5141
3.863 | -98.4455
3.8667 | -98.3626
3.8703 | -98.2944
3.874 | -98.2472
3.8776 | -98.2433
3.8813 | -98.2645
3.885 | -98.3117
3.8886 | -98.3764
3.8923 | -98.4493
3.8959 | -98.528
3.8996 | -98.6086
3.9032 | -98.6836
3.9069 | -98.7489
3.9106 | -98.787
3.9142 | -98.8169
3.9179 | -98.8489
3.9215 | -98.8654
3.9252 | -98.8898
3.9288 | -98.9311
3.9325 | -98.9575
3.9362 | -98.9652
3.9398 | -98.9426
3.9435 | -98.8815
3.9471 | -98.7775
3.9508 | -98.6385
3.9545 | -98.4805
3.9581 | -98.3218
3.9618 | -98.1861
3.9654 | -98.0885
3.9691 | -98.0355
3.9727 | -98.0362
3.9764 | -98.0873
3.9801 | -98.2001
3.9837 | -98.3694
3.9874 | -98.562
3.991 | -98.7482
3.9947 | -98.9127
3.9984 | -99.0414
4.002 | -99.1158
4.0057 | -99.1349
4.0093 | -99.0988
4.013 | -99.034
4.0166 | -98.9496
4.0203 | -98.8623
4.024 | -98.7866
4.0276 | -98.7266
4.0313 | -98.6966
4.0349 | -98.7049
4.0386 | -98.7607
4.0423 | -98.8624
4.0459 | -99.005
4.0496 | -99.1815
4.0532 | -99.3752
4.0569 | -99.5375
4.0605 | -99.6537
4.0642 | -99.7162
4.0679 | -99.7302
4.0715 | -99.7047
4.0752 | -99.6442
4.0788 | -99.5669
4.0825 | -99.4961
4.0861 | -99.436
4.0898 | -99.3984
4.0935 | -99.3801
4.0971 | -99.3724
4.1008 | -99.3632
4.1044 | -99.3574
4.1081 | -99.3562
4.1118 | -99.3283
4.1154 | -99.2705
4.1191 | -99.1949
4.1227 | -99.0945
4.1264 | -98.988
4.13 | -98.8999
4.1337 | -98.861
4.1374 | -98.8857
4.141 | -98.9746
4.1447 | -99.1142
4.1483 | -99.2858
4.152 | -99.4518
4.1557 | -99.5828
4.1593 | -99.6494
4.163 | -99.6499
4.1666 | -99.5946
4.1703 | -99.4982
4.1739 | -99.3808
4.1776 | -99.2646
4.1813 | -99.1683
4.1849 | -99.0968
4.1886 | -99.0604
4.1922 | -99.0556
4.1959 | -99.0607
4.1996 | -99.0689
4.2032 | -99.0703
4.2069 | -99.0498
4.2105 | -99.0204
4.2142 | -98.9979
4.2178 | -98.9965
4.2215 | -99.0635
4.2252 | -99.1252
4.2288 | -99.1971
4.2325 | -99.2764
4.2361 | -99.3665
4.2398 | -99.4588
4.2435 | -99.5253
4.2471 | -99.5675
4.2508 | -99.5846
4.2544 | -99.5801
4.2581 | -99.5501
4.2617 | -99.5088
4.2654 | -99.4587
4.2691 | -99.406
4.2727 | -99.3584
4.2764 | -99.3231
4.28 | -99.2972
4.2837 | -99.2857
4.2873 | -99.2954
4.291 | -99.3109
4.2947 | -99.3349
4.2983 | -99.3773
4.302 | -99.4212
4.3056 | -99.4605
4.3093 | -99.5089
4.313 | -99.5652
4.3166 | -99.6241
4.3203 | -99.692
4.3239 | -99.7511
4.3276 | -99.8048
4.3312 | -99.8407
4.3349 | -99.8558
4.3386 | -99.8521
4.3422 | -99.8313
4.3459 | -99.7994
4.3495 | -99.7664
4.3532 | -99.737
4.3569 | -99.7242
4.3605 | -99.7282
4.3642 | -99.7431
4.3678 | -99.7617
4.3715 | -99.7893
4.3751 | -99.819
4.3788 | -99.8416
4.3825 | -99.8655
4.3861 | -99.8885
4.3898 | -99.906
4.3934 | -99.9235
4.3971 | -99.9356
4.4008 | -99.9379
4.4044 | -99.9252
4.4081 | -99.8943
4.4117 | -99.8494
4.4154 | -99.7958
4.419 | -99.7425
4.4227 | -99.6973
4.4264 | -99.6687
4.43 | -99.6692
4.4337 | -99.6949
4.4373 | -99.7553
4.441 | -99.8351
4.4446 | -99.9229
4.4483 | -100.0048
4.452 | -100.0581
4.4556 | -100.0669
4.4593 | -100.0362
4.4629 | -99.9651
4.4666 | -99.8661
4.4703 | -99.7546
4.4739 | -99.6432
4.4776 | -99.5421
4.4812 | -99.4633
4.4849 | -99.4092
4.4885 | -99.3893
4.4922 | -99.4039
4.4959 | -99.4426
4.4995 | -99.5047
4.5032 | -99.5874
4.5068 | -99.6731
4.5105 | -99.7647
4.5142 | -99.8487
4.5178 | -99.9208
4.5215 | -99.9941
4.5251 | -100.0574
4.5288 | -100.1009
4.5324 | -100.1188
4.5361 | -100.109
4.5398 | -100.0845
4.5434 | -100.0414
4.5471 | -99.9907
4.5507 | -99.9323
4.5544 | -99.8733
4.558 | -99.8251
4.5617 | -99.7871
4.5654 | -99.7778
4.569 | -99.7943
4.5727 | -99.827
4.5763 | -99.8683
4.58 | -99.9141
4.5837 | -99.9692
4.5873 | -100.0175
4.591 | -100.0451
4.5946 | -100.0611
4.5983 | -100.0619
4.6019 | -100.0415
4.6056 | -100.001
4.6093 | -99.9428
4.6129 | -99.8908
4.6166 | -99.858
4.6202 | -99.8534
4.6239 | -99.8873
4.6276 | -99.9454
4.6312 | -100.0237
4.6349 | -100.0879
4.6385 | -100.1294
4.6422 | -100.1516
4.6458 | -100.1369
4.6495 | -100.0809
4.6532 | -99.9964
4.6568 | -99.9099
4.6605 | -99.836
4.6641 | -99.7922
4.6678 | -99.796
4.6715 | -99.851
4.6751 | -99.95
4.6788 | -100.0839
4.6824 | -100.2345
4.6861 | -100.3679
4.6897 | -100.4772
4.6934 | -100.5403
4.6971 | -100.557
4.7007 | -100.5323
4.7044 | -100.481
4.708 | -100.4052
4.7117 | -100.3198
4.7153 | -100.2604
4.719 | -100.2186
4.7227 | -100.1892
4.7263 | -100.1747
4.73 | -100.1667
4.7336 | -100.1538
4.7373 | -100.1364
4.741 | -100.1336
4.7446 | -100.1429
4.7483 | -100.1964
4.7519 | -100.2387
4.7556 | -100.2838
4.7592 | -100.3447
4.7629 | -100.3961
4.7666 | -100.43
4.7702 | -100.4535
4.7739 | -100.4665
4.7775 | -100.4657
4.7812 | -100.4539
4.7849 | -100.4304
4.7885 | -100.3974
4.7922 | -100.3714
4.7958 | -100.3485
4.7995 | -100.333
4.8031 | -100.3431
4.8068 | -100.3777
4.8105 | -100.4248
4.8141 | -100.4836
4.8178 | -100.5523
4.8214 | -100.6247
4.8251 | -100.6905
4.8288 | -100.7416
4.8324 | -100.7724
4.8361 | -100.7651
4.8397 | -100.7323
4.8434 | -100.6741
4.847 | -100.59
4.8507 | -100.495
4.8544 | -100.4137
4.858 | -100.3458
4.8617 | -100.2952
4.8653 | -100.2667
4.869 | -100.2591
4.8727 | -100.2744
4.8763 | -100.2987
4.88 | -100.3311
4.8836 | -100.3739
4.8873 | -100.4204
4.8909 | -100.4618
4.8946 | -100.4971
4.8983 | -100.5186
4.9019 | -100.525
4.9056 | -100.5122
4.9092 | -100.4854
4.9129 | -100.4525
4.9165 | -100.4145
4.9202 | -100.3888
4.9239 | -100.3754
4.9275 | -100.3715
4.9312 | -100.3759
4.9348 | -100.396
4.9385 | -100.4251
4.9422 | -100.4602
4.9458 | -100.4977
4.9495 | -100.5423
4.9531 | -100.6113
4.9568 | -100.6951
4.9604 | -100.7956
4.9641 | -100.911
4.9678 | -101.0289
4.9714 | -101.1411
4.9751 | -101.2324
4.9787 | -101.29
4.9824 | -101.305
4.9861 | -101.2731
4.9897 | -101.19
4.9934 | -101.0624
4.997 | -100.899
5.0007 | -100.7232
5.0043 | -100.5534
5.008 | -100.3958
5.0117 | -100.2575
5.0153 | -100.1527
5.019 | -100.0703
5.0226 | -100.0143
5.0263 | -99.9895
5.03 | -99.9822
5.0336 | -99.987
5.0373 | -100.005
5.0409 | -100.036
5.0446 | -100.0723
5.0482 | -100.103
5.0519 | -100.1268
5.0556 | -100.1417
5.0592 | -100.1488
5.0629 | -100.1466
5.0665 | -100.1465
5.0702 | -100.1468
5.0738 | -100.1538
5.0775 | -100.1696
5.0812 | -100.195
5.0848 | -100.2294
5.0885 | -100.2773
5.0921 | -100.3297
5.0958 | -100.3871
5.0995 | -100.444
5.1031 | -100.4936
5.1068 | -100.5422
5.1104 | -100.5997
5.1141 | -100.6542
5.1177 | -100.7051
5.1214 | -100.7619
5.1251 | -100.8264
5.1287 | -100.8875
5.1324 | -100.9444
5.136 | -100.9944
5.1397 | -101.0343
5.1434 | -101.0577
5.147 | -101.0787
5.1507 | -101.082
5.1543 | -101.074
5.158 | -101.0563
5.1616 | -101.0182
5.1653 | -100.9697
5.169 | -100.9178
5.1726 | -100.8657
5.1763 | -100.8137
5.1799 | -100.7612
5.1836 | -100.7215
5.1873 | -100.6992
5.1909 | -100.6897
5.1946 | -100.6955
5.1982 | -100.7079
5.2019 | -100.7285
5.2055 | -100.7615
5.2092 | -100.7972
5.2129 | -100.822
5.2165 | -100.8391
5.2202 | -100.8444
5.2238 | -100.8305
5.2275 | -100.8142
5.2311 | -100.7903
5.2348 | -100.7707
5.2385 | -100.7702
5.2421 | -100.7716
5.2458 | -100.7811
5.2494 | -100.803
5.2531 | -100.8328
5.2568 | -100.8716
5.2604 | -100.9146
5.2641 | -100.9586
5.2677 | -101.0169
5.2714 | -101.0791
5.275 | -101.121
5.2787 | -101.1601
5.2824 | -101.1852
5.286 | -101.1906
5.2897 | -101.174
5.2933 | -101.1366
5.297 | -101.0795
5.3007 | -101.0177
5.3043 | -100.956
5.308 | -100.9015
5.3116 | -100.8636
5.3153 | -100.842
5.3189 | -100.8414
5.3226 | -100.8627
5.3263 | -100.901
5.3299 | -100.9447
5.3336 | -100.9991
5.3372 | -101.0515
5.3409 | -101.1097
5.3446 | -101.1691
5.3482 | -101.2191
5.3519 | -101.2679
5.3555 | -101.3145
5.3592 | -101.359
5.3628 | -101.3995
5.3665 | -101.4378
5.3702 | -101.4691
5.3738 | -101.4995
5.3775 | -101.522
5.3811 | -101.537
5.3848 | -101.546
5.3884 | -101.5505
5.3921 | -101.5436
5.3958 | -101.5368
5.3994 | -101.5349
5.4031 | -101.5369
5.4067 | -101.5458
5.4104 | -101.5722
5.4141 | -101.6221
5.4177 | -101.6866
5.4214 | -101.7686
5.425 | -101.8513
5.4287 | -101.9329
5.4323 | -102.012
5.436 | -102.0716
5.4397 | -102.1048
5.4433 | -102.1129
5.447 | -102.0983
5.4506 | -102.0582
5.4543 | -102.0015
5.458 | -101.9393
5.4616 | -101.8796
5.4653 | -101.8187
5.4689 | -101.7589
5.4726 | -101.7078
5.4762 | -101.6633
5.4799 | -101.6287
5.4836 | -101.6033
5.4872 | -101.5785
5.4909 | -101.5572
5.4945 | -101.5309
5.4982 | -101.4922
5.5019 | -101.4356
5.5055 | -101.3726
5.5092 | -101.3042
5.5128 | -101.2334
5.5165 | -101.1656
5.5201 | -101.1067
5.5238 | -101.0629
5.5275 | -101.037
5.5311 | -101.0175
5.5348 | -101.0184
5.5384 | -101.0308
5.5421 | -101.0405
5.5457 | -101.0482
5.5494 | -101.0491
5.5531 | -101.0454
5.5567 | -101.0394
5.5604 | -101.0346
5.564 | -101.0292
5.5677 | -101.0293
5.5714 | -101.0505
5.575 | -101.089
5.5787 | -101.1411
5.5823 | -101.206
5.586 | -101.283
5.5896 | -101.3604
5.5933 | -101.4272
5.597 | -101.4824
5.6006 | -101.5173
5.6043 | -101.5369
5.6079 | -101.5378
5.6116 | -101.5242
5.6153 | -101.5151
5.6189 | -101.5034
5.6226 | -101.5003
5.6262 | -101.5077
5.6299 | -101.5256
5.6335 | -101.5476
5.6372 | -101.5706
5.6409 | -101.5907
5.6445 | -101.5992
5.6482 | -101.6034
5.6518 | -101.5926
5.6555 | -101.5574
5.6592 | -101.5047
5.6628 | -101.4459
5.6665 | -101.383
5.6701 | -101.3069
5.6738 | -101.23
5.6774 | -101.1603
5.6811 | -101.1044
5.6848 | -101.0634
5.6884 | -101.0403
5.6921 | -101.0406
5.6957 | -101.0621
5.6994 | -101.1052
5.7031 | -101.1684
5.7067 | -101.2458
5.7104 | -101.3337
5.714 | -101.4257
5.7177 | -101.5188
5.7213 | -101.5994
5.725 | -101.66
5.7287 | -101.6998
5.7323 | -101.7088
5.736 | -101.7013
5.7396 | -101.6817
5.7433 | -101.6511
5.7469 | -101.6198
5.7506 | -101.5964
5.7543 | -101.5916
5.7579 | -101.5999
5.7616 | -101.6303
5.7652 | -101.6704
5.7689 | -101.7204
5.7726 | -101.7743
5.7762 | -101.8282
5.7799 | -101.88
5.7835 | -101.9185
5.7872 | -101.946
5.7908 | -101.9644
5.7945 | -101.9678
5.7982 | -101.9597
5.8018 | -101.9142
5.8055 | -101.8802
5.8091 | -101.8339
5.8128 | -101.7865
5.8165 | -101.7297
5.8201 | -101.6651
5.8238 | -101.6016
5.8274 | -101.5382
5.8311 | -101.4813
5.8347 | -101.4344
5.8384 | -101.4013
5.8421 | -101.3867
5.8457 | -101.3928
5.8494 | -101.4105
5.853 | -101.4449
5.8567 | -101.4967
5.8604 | -101.5591
5.864 | -101.628
5.8677 | -101.7014
5.8713 | -101.779
5.875 | -101.8544
5.8786 | -101.9226
5.8823 | -101.9818
5.886 | -102.0236
5.8896 | -102.0417
5.8933 | -102.0517
5.8969 | -102.039
5.9006 | -102.0093
5.9042 | -101.9568
5.9079 | -101.8926
5.9116 | -101.8176
5.9152 | -101.7243
5.9189 | -101.6297
5.9225 | -101.5355
5.9262 | -101.4441
5.9299 | -101.3682
5.9335 | -101.3034
5.9372 | -101.2537
5.9408 | -101.227
5.9445 | -101.2186
5.9481 | -101.2299
5.9518 | -101.2584
5.9555 | -101.305
5.9591 | -101.3631
5.9628 | -101.4335
5.9664 | -101.5141
5.9701 | -101.606
5.9738 | -101.7009
5.9774 | -101.7959
5.9811 | -101.894
5.9847 | -101.972
5.9884 | -102.0374
5.992 | -102.0871
5.9957 | -102.1218
5.9994 | -102.1414
6.003 | -102.146
6.0067 | -102.1424
6.0103 | -102.127
6.014 | -102.1114
6.0177 | -102.1014
6.0213 | -102.0927
6.025 | -102.0895
6.0286 | -102.0967
6.0323 | -102.1067
6.0359 | -102.1214
6.0396 | -102.1329
6.0433 | -102.1469
6.0469 | -102.1608
6.0506 | -102.1751
6.0542 | -102.181
6.0579 | -102.1836
6.0615 | -102.1832
6.0652 | -102.1838
6.0689 | -102.1743
6.0725 | -102.157
6.0762 | -102.136
6.0798 | -102.1054
6.0835 | -102.0743
6.0872 | -102.0343
6.0908 | -101.9881
6.0945 | -101.9489
6.0981 | -101.9206
6.1018 | -101.8978
6.1054 | -101.877
6.1091 | -101.8591
6.1128 | -101.8477
6.1164 | -101.8429
6.1201 | -101.8409
6.1237 | -101.8416
6.1274 | -101.8496
6.1311 | -101.859
6.1347 | -101.8766
6.1384 | -101.8936
6.142 | -101.9068
6.1457 | -101.9197
6.1493 | -101.9218
6.153 | -101.9071
6.1567 | -101.8855
6.1603 | -101.864
6.164 | -101.839
6.1676 | -101.806
6.1713 | -101.7733
6.175 | -101.7507
6.1786 | -101.7494
6.1823 | -101.7627
6.1859 | -101.7845
6.1896 | -101.8141
6.1932 | -101.8572
6.1969 | -101.9122
6.2006 | -101.9733
6.2042 | -102.0392
6.2079 | -102.1078
6.2115 | -102.17
6.2152 | -102.2225
6.2188 | -102.2661
6.2225 | -102.3082
6.2262 | -102.3356
6.2298 | -102.3554
6.2335 | -102.3711
6.2371 | -102.372
6.2408 | -102.365
6.2445 | -102.3469
6.2481 | -102.3243
6.2518 | -102.2954
6.2554 | -102.256
6.2591 | -102.2233
6.2627 | -102.1873
6.2664 | -102.1408
6.2701 | -102.0968
6.2737 | -102.0621
6.2774 | -102.0336
6.281 | -102.0175
6.2847 | -102.009
6.2884 | -102.0024
6.292 | -102.0055
6.2957 | -102.0184
6.2993 | -102.0409
6.303 | -102.0707
6.3066 | -102.1151
6.3103 | -102.1594
6.314 | -102.1945
6.3176 | -102.2229
6.3213 | -102.2496
6.3249 | -102.2626
6.3286 | -102.2746
6.3323 | -102.257
6.3359 | -102.2526
6.3396 | -102.2402
6.3432 | -102.2115
6.3469 | -102.1788
6.3505 | -102.1512
6.3542 | -102.1162
6.3579 | -102.0858
6.3615 | -102.0566
6.3652 | -102.0293
6.3688 | -102.0119
6.3725 | -102.0034
6.3761 | -101.9976
6.3798 | -101.9998
6.3835 | -102.0126
6.3871 | -102.0256
6.3908 | -102.0426
6.3944 | -102.0634
6.3981 | -102.0917
6.4018 | -102.1195
6.4054 | -102.1509
6.4091 | -102.1884
6.4127 | -102.2257
6.4164 | -102.2612
6.42 | -102.3015
6.4237 | -102.3399
6.4274 | -102.3787
6.431 | -102.4234
6.4347 | -102.4669
6.4383 | -102.514
6.442 | -102.5586
6.4457 | -102.5968
6.4493 | -102.6263
6.453 | -102.65
6.4566 | -102.6607
6.4603 | -102.6654
6.4639 | -102.6586
6.4676 | -102.6423
6.4713 | -102.6238
6.4749 | -102.6001
6.4786 | -102.5711
6.4822 | -102.5427
6.4859 | -102.5163
6.4896 | -102.4853
6.4932 | -102.4517
6.4969 | -102.4174
6.5005 | -102.3862
6.5042 | -102.3624
6.5078 | -102.3431
6.5115 | -102.3276
6.5152 | -102.3132
6.5188 | -102.3017
6.5225 | -102.2911
6.5261 | -102.2916
6.5298 | -102.3015
6.5334 | -102.3157
6.5371 | -102.3359
6.5408 | -102.3615
6.5444 | -102.3864
6.5481 | -102.4099
6.5517 | -102.4341
6.5554 | -102.4534
6.5591 | -102.4688
6.5627 | -102.4801
6.5664 | -102.4868
6.57 | -102.4919
6.5737 | -102.5007
6.5773 | -102.5086
6.581 | -102.5147
6.5847 | -102.5205
6.5883 | -102.53
6.592 | -102.5464
6.5956 | -102.5601
6.5993 | -102.577
6.603 | -102.5874
6.6066 | -102.6037
6.6103 | -102.6249
6.6139 | -102.6468
6.6176 | -102.6642
6.6212 | -102.6866
6.6249 | -102.7095
6.6286 | -102.7342
6.6322 | -102.7622
6.6359 | -102.7823
6.6395 | -102.8036
6.6432 | -102.8246
6.6469 | -102.8451
6.6505 | -102.8594
6.6542 | -102.8731
6.6578 | -102.8808
6.6615 | -102.8789
6.6651 | -102.8636
6.6688 | -102.8408
6.6725 | -102.8085
6.6761 | -102.7644
6.6798 | -102.7116
6.6834 | -102.6558
6.6871 | -102.5973
6.6907 | -102.5343
6.6944 | -102.4792
6.6981 | -102.4249
6.7017 | -102.3838
6.7054 | -102.3491
6.709 | -102.3248
6.7127 | -102.3211
6.7164 | -102.3374
6.72 | -102.3632
6.7237 | -102.3951
6.7273 | -102.4376
6.731 | -102.4879
6.7346 | -102.5423
6.7383 | -102.5886
6.742 | -102.6296
6.7456 | -102.6629
6.7493 | -102.6892
6.7529 | -102.7069
6.7566 | -102.7288
6.7603 | -102.7513
6.7639 | -102.7777
6.7676 | -102.8035
6.7712 | -102.8348
6.7749 | -102.8752
6.7785 | -102.9129
6.7822 | -102.9432
6.7859 | -102.9755
6.7895 | -103.0088
6.7932 | -103.0411
6.7968 | -103.0731
6.8005 | -103.1009
6.8042 | -103.1256
6.8078 | -103.1469
6.8115 | -103.1573
6.8151 | -103.1586
6.8188 | -103.1597
6.8224 | -103.1558
6.8261 | -103.1446
6.8298 | -103.1343
6.8334 | -103.1221
6.8371 | -103.112
6.8407 | -103.1035
6.8444 | -103.1041
6.848 | -103.1134
6.8517 | -103.1251
6.8554 | -103.1252
6.859 | -103.1215
6.8627 | -103.1201
6.8663 | -103.1106
6.87 | -103.0947
6.8737 | -103.078
6.8773 | -103.0624
6.881 | -103.0489
6.8846 | -103.0411
6.8883 | -103.0341
6.8919 | -103.0311
6.8956 | -103.0358
6.8993 | -103.0478
6.9029 | -103.0663
6.9066 | -103.0931
6.9102 | -103.1188
6.9139 | -103.1481
6.9176 | -103.1773
6.9212 | -103.2023
6.9249 | -103.2271
6.9285 | -103.245
6.9322 | -103.2535
6.9358 | -103.2613
6.9395 | -103.255
6.9432 | -103.2431
6.9468 | -103.2228
6.9505 | -103.195
6.9541 | -103.1711
6.9578 | -103.147
6.9615 | -103.1183
6.9651 | -103.089
6.9688 | -103.0655
6.9724 | -103.0408
6.9761 | -103.0128
6.9797 | -102.9809
6.9834 | -102.9412
6.9871 | -102.9036
6.9907 | -102.8682
6.9944 | -102.8303
6.998 | -102.7901
7.0017 | -102.7446
7.0053 | -102.7012
7.009 | -102.6593
7.0127 | -102.6176
7.0163 | -102.5816
7.02 | -102.5443
7.0236 | -102.5135
7.0273 | -102.4939
7.031 | -102.4854
7.0346 | -102.487
7.0383 | -102.4976
7.0419 | -102.5159
7.0456 | -102.5464
7.0492 | -102.5792
7.0529 | -102.609
7.0566 | -102.6382
7.0602 | -102.6665
7.0639 | -102.6993
7.0675 | -102.7269
7.0712 | -102.7448
7.0749 | -102.7583
7.0785 | -102.7709
7.0822 | -102.7804
7.0858 | -102.7856
7.0895 | -102.7944
7.0931 | -102.8013
7.0968 | -102.8023
7.1005 | -102.8079
7.1041 | -102.8244
7.1078 | -102.8479
7.1114 | -102.8722
7.1151 | -102.9049
7.1188 | -102.9439
7.1224 | -102.9882
7.1261 | -103.0342
7.1297 | -103.0718
7.1334 | -103.0937
7.137 | -103.1094
7.1407 | -103.1192
7.1444 | -103.1218
7.148 | -103.1123
7.1517 | -103.0931
7.1553 | -103.0618
7.159 | -103.0268
7.1627 | -102.987
7.1663 | -102.9406
7.17 | -102.8865
7.1736 | -102.8411
7.1773 | -102.8075
7.1809 | -102.7812
7.1846 | -102.7637
7.1883 | -102.7531
7.1919 | -102.7506
7.1956 | -102.7557
7.1992 | -102.7645
7.2029 | -102.781
7.2065 | -102.7995
7.2102 | -102.8134
7.2139 | -102.8232
7.2175 | -102.8326
7.2212 | -102.8398
7.2248 | -102.8427
7.2285 | -102.8378
7.2322 | -102.8312
7.2358 | -102.8253
7.2395 | -102.8221
7.2431 | -102.8173
7.2468 | -102.8147
7.2504 | -102.8106
7.2541 | -102.8104
7.2578 | -102.8151
7.2614 | -102.8284
7.2651 | -102.8485
7.2687 | -102.8737
7.2724 | -102.8952
7.2761 | -102.9259
7.2797 | -102.96
7.2834 | -102.991
7.287 | -103.022
7.2907 | -103.054
7.2943 | -103.0866
7.298 | -103.1178
7.3017 | -103.1495
7.3053 | -103.1784
7.309 | -103.1977
7.3126 | -103.2109
7.3163 | -103.2229
7.32 | -103.2257
7.3236 | -103.2158
7.3273 | -103.1927
7.3309 | -103.1672
7.3346 | -103.1373
7.3382 | -103.0964
7.3419 | -103.0543
7.3456 | -103.0082
7.3492 | -102.9592
7.3529 | -102.9084
7.3565 | -102.8609
7.3602 | -102.8162
7.3638 | -102.7759
7.3675 | -102.7404
7.3712 | -102.7096
7.3748 | -102.6849
7.3785 | -102.6752
7.3821 | -102.6672
7.3858 | -102.675
7.3895 | -102.6715
7.3931 | -102.6741
7.3968 | -102.6833
7.4004 | -102.6934
7.4041 | -102.7013
7.4077 | -102.7112
7.4114 | -102.7175
7.4151 | -102.7264
7.4187 | -102.7344
7.4224 | -102.7445
7.426 | -102.7574
7.4297 | -102.7711
7.4334 | -102.7836
7.437 | -102.8001
7.4407 | -102.8158
7.4443 | -102.8317
7.448 | -102.8485
7.4516 | -102.8577
7.4553 | -102.8645
7.459 | -102.8715
7.4626 | -102.8799
7.4663 | -102.89
7.4699 | -102.8986
7.4736 | -102.9071
7.4773 | -102.9206
7.4809 | -102.9396
7.4846 | -102.9627
7.4882 | -102.9885
7.4919 | -103.018
7.4955 | -103.0483
7.4992 | -103.073
7.5029 | -103.1009
7.5065 | -103.1255
7.5102 | -103.1407
7.5138 | -103.1534
7.5175 | -103.1637
7.5211 | -103.1616
7.5248 | -103.1547
7.5285 | -103.1397
7.5321 | -103.1234
7.5358 | -103.1007
7.5394 | -103.0748
7.5431 | -103.0556
7.5468 | -103.0419
7.5504 | -103.0297
7.5541 | -103.0233
7.5577 | -103.029
7.5614 | -103.0421
7.565 | -103.0686
7.5687 | -103.1068
7.5724 | -103.1525
7.576 | -103.2057
7.5797 | -103.2658
7.5833 | -103.3286
7.587 | -103.3913
7.5907 | -103.4556
7.5943 | -103.5136
7.598 | -103.5672
7.6016 | -103.6168
7.6053 | -103.6616
7.6089 | -103.6947
7.6126 | -103.7157
7.6163 | -103.7284
7.6199 | -103.7335
7.6236 | -103.7324
7.6272 | -103.7246
7.6309 | -103.7065
7.6346 | -103.6833
7.6382 | -103.6577
7.6419 | -103.6329
7.6455 | -103.6021
7.6492 | -103.5666
7.6528 | -103.5355
7.6565 | -103.5032
7.6602 | -103.4724
7.6638 | -103.4507
7.6675 | -103.4388
7.6711 | -103.4386
7.6748 | -103.4512
7.6784 | -103.4726
7.6821 | -103.5041
7.6858 | -103.5474
7.6894 | -103.5988
7.6931 | -103.6593
7.6967 | -103.723
7.7004 | -103.7809
7.7041 | -103.8349
7.7077 | -103.8826
7.7114 | -103.9213
7.715 | -103.9494
7.7187 | -103.9641
7.7223 | -103.9676
7.726 | -103.9555
7.7297 | -103.928
7.7333 | -103.8881
7.737 | -103.8372
7.7406 | -103.7823
7.7443 | -103.7258
7.748 | -103.6694
7.7516 | -103.6127
7.7553 | -103.5591
7.7589 | -103.5142
7.7626 | -103.4808
7.7662 | -103.4544
7.7699 | -103.4412
7.7736 | -103.4396
7.7772 | -103.4402
7.7809 | -103.4493
7.7845 | -103.4645
7.7882 | -103.4878
7.7919 | -103.5185
7.7955 | -103.5508
7.7992 | -103.5812
7.8028 | -103.6166
7.8065 | -103.6534
7.8101 | -103.6949
7.8138 | -103.7336
7.8175 | -103.7598
7.8211 | -103.7818
7.8248 | -103.7915
7.8284 | -103.7845
7.8321 | -103.7722
7.8357 | -103.7535
7.8394 | -103.7306
7.8431 | -103.6946
7.8467 | -103.658
7.8504 | -103.6159
7.854 | -103.566
7.8577 | -103.5182
7.8614 | -103.4746
7.865 | -103.4334
7.8687 | -103.3959
7.8723 | -103.3625
7.876 | -103.3346
7.8796 | -103.3167
7.8833 | -103.3008
7.887 | -103.2875
7.8906 | -103.2845
7.8943 | -103.2854
7.8979 | -103.2904
7.9016 | -103.299
7.9053 | -103.3105
7.9089 | -103.3213
7.9126 | -103.3368
7.9162 | -103.3477
7.9199 | -103.3615
7.9235 | -103.3713
7.9272 | -103.3802
7.9309 | -103.3929
7.9345 | -103.4091
7.9382 | -103.4236
7.9418 | -103.4365
7.9455 | -103.4534
7.9492 | -103.4672
7.9528 | -103.4837
7.9565 | -103.5006
7.9601 | -103.522
7.9638 | -103.5455
7.9674 | -103.5733
7.9711 | -103.6018
7.9748 | -103.6278
7.9784 | -103.653
7.9821 | -103.6772
7.9857 | -103.7005
7.9894 | -103.72
7.9931 | -103.7345
7.9967 | -103.7387
8.0004 | -103.7465
8.004 | -103.751
8.0077 | -103.7448
8.0113 | -103.7355
8.015 | -103.7208
8.0187 | -103.7084
8.0223 | -103.6936
8.026 | -103.6848
8.0296 | -103.6736
8.0333 | -103.6637
8.0369 | -103.6553
8.0406 | -103.6481
8.0443 | -103.6414
8.0479 | -103.6384
8.0516 | -103.6368
8.0552 | -103.6381
8.0589 | -103.6449
8.0626 | -103.6525
8.0662 | -103.6584
8.0699 | -103.6634
8.0735 | -103.6691
8.0772 | -103.6785
8.0808 | -103.689
8.0845 | -103.6951
8.0882 | -103.7021
8.0918 | -103.707
8.0955 | -103.7097
8.0991 | -103.7117
8.1028 | -103.7049
8.1065 | -103.6996
8.1101 | -103.6986
8.1138 | -103.6954
8.1174 | -103.6977
8.1211 | -103.6998
8.1247 | -103.7045
8.1284 | -103.7136
8.1321 | -103.7187
8.1357 | -103.7297
8.1394 | -103.7398
8.143 | -103.7499
8.1467 | -103.7624
8.1503 | -103.7743
8.154 | -103.7805
8.1577 | -103.7892
8.1613 | -103.8024
8.165 | -103.817
8.1686 | -103.8248
8.1723 | -103.8333
8.176 | -103.8422
8.1796 | -103.8521
8.1833 | -103.8662
8.1869 | -103.877
8.1906 | -103.8827
8.1942 | -103.8848
8.1979 | -103.8893
8.2016 | -103.891
8.2052 | -103.8982
8.2089 | -103.9076
8.2125 | -103.9132
8.2162 | -103.9195
8.2199 | -103.921
8.2235 | -103.9256
8.2272 | -103.9275
8.2308 | -103.9244
8.2345 | -103.9264
8.2381 | -103.9232
8.2418 | -103.9186
8.2455 | -103.9197
8.2491 | -103.9225
8.2528 | -103.9297
8.2564 | -103.9352
8.2601 | -103.9399
8.2638 | -103.9481
8.2674 | -103.9524
8.2711 | -103.9562
8.2747 | -103.9621
8.2784 | -103.9625
8.282 | -103.9538
8.2857 | -103.9421
8.2894 | -103.929
8.293 | -103.9091
8.2967 | -103.8818
8.3003 | -103.8504
8.304 | -103.8227
8.3077 | -103.7972
8.3113 | -103.7782
8.315 | -103.7574
8.3186 | -103.7357
8.3223 | -103.7251
8.3259 | -103.7154
8.3296 | -103.7013
8.3333 | -103.6924
8.3369 | -103.6834
8.3406 | -103.6738
8.3442 | -103.6642
8.3479 | -103.6538
8.3515 | -103.6446
8.3552 | -103.6284
8.3589 | -103.6145
8.3625 | -103.6023
8.3662 | -103.5861
8.3698 | -103.5744
8.3735 | -103.5661
8.3772 | -103.5611
8.3808 | -103.5617
8.3845 | -103.5648
8.3881 | -103.5743
8.3918 | -103.5933
8.3954 | -103.6129
8.3991 | -103.6339
8.4028 | -103.6614
8.4064 | -103.6918
8.4101 | -103.7234
8.4137 | -103.7542
8.4174 | -103.7877
8.4211 | -103.8169
8.4247 | -103.8417
8.4284 | -103.8687
8.432 | -103.8918
8.4357 | -103.9145
8.4393 | -103.9334
8.443 | -103.9396
8.4467 | -103.9498
8.4503 | -103.9549
8.454 | -103.9585
8.4576 | -103.9559
8.4613 | -103.9465
8.4649 | -103.934
8.4686 | -103.9195
8.4723 | -103.8968
8.4759 | -103.8674
8.4796 | -103.8414
8.4832 | -103.816
8.4869 | -103.7881
8.4906 | -103.7631
8.4942 | -103.7358
8.4979 | -103.7067
8.5015 | -103.6824
8.5052 | -103.6663
8.5088 | -103.6581
8.5125 | -103.658
8.5162 | -103.659
8.5198 | -103.6663
8.5235 | -103.6777
8.5271 | -103.6899
8.5308 | -103.7041
8.5345 | -103.7213
8.5381 | -103.74
8.5418 | -103.7549
8.5454 | -103.7698
8.5491 | -103.7811
8.5527 | -103.7885
8.5564 | -103.7942
8.5601 | -103.7948
8.5637 | -103.7945
8.5674 | -103.7975
8.571 | -103.7984
8.5747 | -103.7993
8.5784 | -103.804
8.582 | -103.8101
8.5857 | -103.8163
8.5893 | -103.8231
8.593 | -103.8251
8.5966 | -103.8266
8.6003 | -103.8293
8.604 | -103.8323
8.6076 | -103.8321
8.6113 | -103.8299
8.6149 | -103.8291
8.6186 | -103.8339
8.6223 | -103.8428
8.6259 | -103.853
8.6296 | -103.8615
8.6332 | -103.8668
8.6369 | -103.8751
8.6405 | -103.8833
8.6442 | -103.8937
8.6479 | -103.9091
8.6515 | -103.927
8.6552 | -103.945
8.6588 | -103.9633
8.6625 | -103.9754
8.6661 | -103.9878
8.6698 | -104.0027
8.6735 | -104.0174
8.6771 | -104.0332
8.6808 | -104.0454
8.6844 | -104.0562
8.6881 | -104.0656
8.6918 | -104.0731
8.6954 | -104.0761
8.6991 | -104.075
8.7027 | -104.0748
8.7064 | -104.0731
8.71 | -104.0674
8.7137 | -104.062
8.7174 | -104.0561
8.721 | -104.0537
8.7247 | -104.0457
8.7283 | -104.037
8.732 | -104.03
8.7357 | -104.0222
8.7393 | -104.0154
8.743 | -104.0121
8.7466 | -104.0076
8.7503 | -104.0074
8.7539 | -104.0103
8.7576 | -104.0163
8.7613 | -104.026
8.7649 | -104.0369
8.7686 | -104.0505
8.7722 | -104.0655
8.7759 | -104.079
8.7796 | -104.092
8.7832 | -104.1058
8.7869 | -104.1216
8.7905 | -104.1399
8.7942 | -104.1534
8.7978 | -104.1651
8.8015 | -104.1764
8.8052 | -104.1868
8.8088 | -104.203
8.8125 | -104.2231
8.8161 | -104.2437
8.8198 | -104.2626
8.8234 | -104.2799
8.8271 | -104.2997
8.8308 | -104.3188
8.8344 | -104.3375
8.8381 | -104.3549
8.8417 | -104.3665
8.8454 | -104.3743
8.8491 | -104.3875
8.8527 | -104.4018
8.8564 | -104.4183
8.86 | -104.4339
8.8637 | -104.4473
8.8673 | -104.4597
8.871 | -104.4707
8.8747 | -104.4817
8.8783 | -104.4932
8.882 | -104.5043
8.8856 | -104.5114
8.8893 | -104.5167
8.893 | -104.5175
8.8966 | -104.5158
8.9003 | -104.5114
8.9039 | -104.503
8.9076 | -104.4939
8.9112 | -104.4826
8.9149 | -104.4699
8.9186 | -104.4549
8.9222 | -104.4417
8.9259 | -104.4312
8.9295 | -104.4207
8.9332 | -104.4076
8.9369 | -104.3901
8.9405 | -104.37
8.9442 | -104.3526
8.9478 | -104.3336
8.9515 | -104.3158
8.9551 | -104.2997
8.9588 | -104.2841
8.9625 | -104.2688
8.9661 | -104.2563
8.9698 | -104.2541
8.9734 | -104.2499
8.9771 | -104.2431
8.9807 | -104.2328
8.9844 | -104.2254
8.9881 | -104.2155
8.9917 | -104.21
8.9954 | -104.2101
8.999 | -104.2121
9.0027 | -104.2136
9.0064 | -104.2115
9.01 | -104.2076
9.0137 | -104.2043
9.0173 | -104.1987
9.021 | -104.1876
9.0246 | -104.1776
9.0283 | -104.1648
9.032 | -104.1513
9.0356 | -104.1347
9.0393 | -104.11
9.0429 | -104.086
9.0466 | -104.0587
9.0503 | -104.034
9.0539 | -104.0089
9.0576 | -103.9868
9.0612 | -103.9697
9.0649 | -103.9591
9.0685 | -103.9493
9.0722 | -103.9423
9.0759 | -103.9351
9.0795 | -103.9325
9.0832 | -103.9352
9.0868 | -103.9398
9.0905 | -103.9452
9.0942 | -103.9536
9.0978 | -103.9636
9.1015 | -103.9758
9.1051 | -103.9893
9.1088 | -104.007
9.1124 | -104.0309
9.1161 | -104.054
9.1198 | -104.076
9.1234 | -104.1029
9.1271 | -104.1272
9.1307 | -104.148
9.1344 | -104.1687
9.138 | -104.192
9.1417 | -104.2174
9.1454 | -104.2424
9.149 | -104.2717
9.1527 | -104.3015
9.1563 | -104.3282
9.16 | -104.3527
9.1637 | -104.3789
9.1673 | -104.4077
9.171 | -104.4374
9.1746 | -104.4688
9.1783 | -104.493
9.1819 | -104.5136
9.1856 | -104.5362
9.1893 | -104.5579
9.1929 | -104.5735
9.1966 | -104.5874
9.2002 | -104.5934
9.2039 | -104.6057
9.2076 | -104.6165
9.2112 | -104.6249
9.2149 | -104.6316
9.2185 | -104.6298
9.2222 | -104.6253
9.2258 | -104.6222
9.2295 | -104.6223
9.2332 | -104.6226
9.2368 | -104.6175
9.2405 | -104.6125
9.2441 | -104.6068
9.2478 | -104.6017
9.2515 | -104.5976
9.2551 | -104.5982
9.2588 | -104.5954
9.2624 | -104.593
9.2661 | -104.5881
9.2697 | -104.5833
9.2734 | -104.579
9.2771 | -104.5726
9.2807 | -104.5627
9.2844 | -104.549
9.288 | -104.5385
9.2917 | -104.5317
9.2953 | -104.5211
9.299 | -104.5046
9.3027 | -104.4829
9.3063 | -104.4658
9.31 | -104.4517
9.3136 | -104.4376
9.3173 | -104.4219
9.321 | -104.406
9.3246 | -104.3891
9.3283 | -104.3717
9.3319 | -104.3597
9.3356 | -104.3559
9.3392 | -104.3484
9.3429 | -104.3452
9.3466 | -104.3446
9.3502 | -104.3436
9.3539 | -104.3491
9.3575 | -104.3585
9.3612 | -104.3667
9.3649 | -104.374
9.3685 | -104.3823
9.3722 | -104.3938
9.3758 | -104.4053
9.3795 | -104.4153
9.3831 | -104.4259
9.3868 | -104.437
9.3905 | -104.4458
9.3941 | -104.4567
9.3978 | -104.4667
9.4014 | -104.4754
9.4051 | -104.4843
9.4088 | -104.49
9.4124 | -104.4982
9.4161 | -104.5034
9.4197 | -104.5042
9.4234 | -104.5024
9.427 | -104.4956
9.4307 | -104.4854
9.4344 | -104.477
9.438 | -104.4686
9.4417 | -104.4528
9.4453 | -104.4379
9.449 | -104.4228
9.4527 | -104.4076
9.4563 | -104.3919
9.46 | -104.3822
9.4636 | -104.3737
9.4673 | -104.3632
9.4709 | -104.3548
9.4746 | -104.351
9.4783 | -104.3517
9.4819 | -104.3539
9.4856 | -104.36
9.4892 | -104.3671
9.4929 | -104.3717
9.4965 | -104.3796
9.5002 | -104.3903
9.5039 | -104.3989
9.5075 | -104.408
9.5112 | -104.4162
9.5148 | -104.4227
9.5185 | -104.43
9.5222 | -104.4341
9.5258 | -104.4363
9.5295 | -104.4387
9.5331 | -104.4367
9.5368 | -104.4379
9.5404 | -104.4436
9.5441 | -104.4468
9.5478 | -104.4472
9.5514 | -104.4407
9.5551 | -104.4347
9.5587 | -104.4315
9.5624 | -104.4279
9.5661 | -104.4276
9.5697 | -104.4296
9.5734 | -104.4269
9.577 | -104.4235
9.5807 | -104.4241
9.5843 | -104.4255
9.588 | -104.4328
9.5917 | -104.4453
9.5953 | -104.4591
9.599 | -104.4753
9.6026 | -104.4969
9.6063 | -104.5219
9.6099 | -104.5513
9.6136 | -104.583
9.6173 | -104.615
9.6209 | -104.6529
9.6246 | -104.6908
9.6282 | -104.723
9.6319 | -104.7591
9.6356 | -104.7919
9.6392 | -104.817
9.6429 | -104.8413
9.6465 | -104.867
9.6502 | -104.887
9.6538 | -104.9023
9.6575 | -104.9118
9.6612 | -104.9216
9.6648 | -104.9286
9.6685 | -104.9313
9.6721 | -104.928
9.6758 | -104.9238
9.6795 | -104.9137
9.6831 | -104.8983
9.6868 | -104.8811
9.6904 | -104.8638
9.6941 | -104.8436
9.6977 | -104.8229
9.7014 | -104.7992
9.7051 | -104.7738
9.7087 | -104.75
9.7124 | -104.723
9.716 | -104.697
9.7197 | -104.6701
9.7234 | -104.6435
9.727 | -104.618
9.7307 | -104.593
9.7343 | -104.5675
9.738 | -104.5453
9.7416 | -104.526
9.7453 | -104.5101
9.749 | -104.4967
9.7526 | -104.4842
9.7563 | -104.477
9.7599 | -104.4698
9.7636 | -104.4657
9.7673 | -104.4657
9.7709 | -104.4733
9.7746 | -104.4793
9.7782 | -104.4902
9.7819 | -104.5039
9.7855 | -104.5174
9.7892 | -104.5347
9.7929 | -104.5543
9.7965 | -104.576
9.8002 | -104.5978
9.8038 | -104.6136
9.8075 | -104.6307
9.8111 | -104.6538
9.8148 | -104.6766
9.8185 | -104.699
9.8221 | -104.7224
9.8258 | -104.7472
9.8294 | -104.77
9.8331 | -104.7947
9.8368 | -104.8234
9.8404 | -104.8492
9.8441 | -104.8683
9.8477 | -104.886
9.8514 | -104.9026
9.855 | -104.9228
9.8587 | -104.9403
9.8624 | -104.955
9.866 | -104.9643
9.8697 | -104.9679
9.8733 | -104.968
9.877 | -104.9684
9.8807 | -104.9597
9.8843 | -104.9471
9.888 | -104.9304
9.8916 | -104.9117
9.8953 | -104.891
9.8989 | -104.8676
9.9026 | -104.8432
9.9063 | -104.8126
9.9099 | -104.7801
9.9136 | -104.7477
9.9172 | -104.7163
9.9209 | -104.6855
9.9246 | -104.6576
9.9282 | -104.6321
9.9319 | -104.6087
9.9355 | -104.591
9.9392 | -104.575
9.9428 | -104.5687
9.9465 | -104.5712
9.9502 | -104.5777
9.9538 | -104.5823
9.9575 | -104.5935
9.9611 | -104.6144
9.9648 | -104.6375
9.9684 | -104.6619
9.9721 | -104.6959
9.9758 | -104.7282
9.9794 | -104.7608
9.9831 | -104.7889
9.9867 | -104.8184
9.9904 | -104.8419
9.9941 | -104.8618
9.9977 | -104.8784
10.0014 | -104.8887
10.005 | -104.8941
10.0087 | -104.8953
10.0123 | -104.8931
10.016 | -104.8874
10.0197 | -104.8691
10.0233 | -104.845
10.027 | -104.8053
10.0306 | -104.7823
10.0343 | -104.7586
10.038 | -104.732
10.0416 | -104.7018
10.0453 | -104.6674
10.0489 | -104.6359
10.0526 | -104.6063
10.0562 | -104.5764
10.0599 | -104.5491
10.0636 | -104.5207
10.0672 | -104.492
10.0709 | -104.4667
10.0745 | -104.4465
10.0782 | -104.4316
10.0819 | -104.4202
10.0855 | -104.414
10.0892 | -104.4067
10.0928 | -104.4023
10.0965 | -104.4053
10.1001 | -104.4149
10.1038 | -104.4282
10.1075 | -104.4446
10.1111 | -104.4647
10.1148 | -104.4865
10.1184 | -104.506
10.1221 | -104.5295
10.1257 | -104.5528
10.1294 | -104.5747
10.1331 | -104.5944
10.1367 | -104.6158
10.1404 | -104.6344
10.144 | -104.6441
10.1477 | -104.6534
10.1514 | -104.6632
10.155 | -104.6704
10.1587 | -104.675
10.1623 | -104.6788
10.166 | -104.6863
10.1696 | -104.6976
10.1733 | -104.7053
10.177 | -104.7084
10.1806 | -104.7116
10.1843 | -104.7165
10.1879 | -104.7232
10.1916 | -104.7376
10.1953 | -104.7556
10.1989 | -104.7746
10.2026 | -104.7949
10.2062 | -104.8159
10.2099 | -104.8347
10.2135 | -104.8553
10.2172 | -104.8737
10.2209 | -104.8891
10.2245 | -104.8997
10.2282 | -104.9061
10.2318 | -104.9111
10.2355 | -104.9109
10.2392 | -104.9089
10.2428 | -104.8994
10.2465 | -104.8875
10.2501 | -104.8789
10.2538 | -104.8696
10.2574 | -104.8576
10.2611 | -104.8483
10.2648 | -104.8417
10.2684 | -104.8342
10.2721 | -104.8282
10.2757 | -104.8231
10.2794 | -104.8227
10.283 | -104.8204
10.2867 | -104.8144
10.2904 | -104.8115
10.294 | -104.8147
10.2977 | -104.8173
10.3013 | -104.8183
10.305 | -104.8207
10.3087 | -104.8237
10.3123 | -104.8242
10.316 | -104.8254
10.3196 | -104.8246
10.3233 | -104.8226
10.3269 | -104.817
10.3306 | -104.8114
10.3343 | -104.8036
10.3379 | -104.7938
10.3416 | -104.7839
10.3452 | -104.7741
10.3489 | -104.7636
10.3526 | -104.7552
10.3562 | -104.749
10.3599 | -104.7418
10.3635 | -104.7402
10.3672 | -104.7441
10.3708 | -104.7527
10.3745 | -104.7641
10.3782 | -104.7774
10.3818 | -104.7945
10.3855 | -104.8156
10.3891 | -104.8331
10.3928 | -104.853
10.3965 | -104.8683
10.4001 | -104.8768
10.4038 | -104.8845
10.4074 | -104.8944
10.4111 | -104.9048
10.4147 | -104.9127
10.4184 | -104.9199
10.4221 | -104.9226
10.4257 | -104.9207
10.4294 | -104.9142
10.433 | -104.9078
10.4367 | -104.9036
10.4403 | -104.8918
10.444 | -104.8745
10.4477 | -104.8549
10.4513 | -104.8338
10.455 | -104.8119
10.4586 | -104.7886
10.4623 | -104.7656
10.466 | -104.741
10.4696 | -104.7145
10.4733 | -104.6884
10.4769 | -104.6645
10.4806 | -104.6406
10.4842 | -104.6186
10.4879 | -104.5996
10.4916 | -104.5792
10.4952 | -104.5658
10.4989 | -104.5569
10.5025 | -104.5531
10.5062 | -104.5527
10.5099 | -104.5596
10.5135 | -104.5681
10.5172 | -104.5815
10.5208 | -104.5983
10.5245 | -104.6216
10.5281 | -104.651
10.5318 | -104.6832
10.5355 | -104.7161
10.5391 | -104.7516
10.5428 | -104.7885
10.5464 | -104.8265
10.5501 | -104.8677
10.5538 | -104.8992
10.5574 | -104.9349
10.5611 | -104.9636
10.5647 | -104.9888
10.5684 | -105.0147
10.572 | -105.0377
10.5757 | -105.0595
10.5794 | -105.0798
10.583 | -105.0953
10.5867 | -105.1069
10.5903 | -105.1142
10.594 | -105.1228
10.5976 | -105.1334
10.6013 | -105.1425
10.605 | -105.1504
10.6086 | -105.1562
10.6123 | -105.1614
10.6159 | -105.1651
10.6196 | -105.1696
10.6233 | -105.1786
10.6269 | -105.1893
10.6306 | -105.2019
10.6342 | -105.2137
10.6379 | -105.2291
10.6415 | -105.2433
10.6452 | -105.2597
10.6489 | -105.2761
10.6525 | -105.2931
10.6562 | -105.3066
10.6598 | -105.3236
10.6635 | -105.3409
10.6672 | -105.3517
10.6708 | -105.3579
10.6745 | -105.3625
10.6781 | -105.3612
10.6818 | -105.3614
10.6854 | -105.3576
10.6891 | -105.3571
10.6928 | -105.3559
10.6964 | -105.3473
10.7001 | -105.3351
10.7037 | -105.3214
10.7074 | -105.3067
10.7111 | -105.2914
10.7147 | -105.2749
10.7184 | -105.2617
10.722 | -105.2476
10.7257 | -105.2328
10.7293 | -105.2174
10.733 | -105.2013
10.7367 | -105.189
10.7403 | -105.1802
10.744 | -105.1708
10.7476 | -105.1614
10.7513 | -105.1543
10.7549 | -105.1492
10.7586 | -105.1458
10.7623 | -105.1419
10.7659 | -105.1451
10.7696 | -105.1525
10.7732 | -105.1637
10.7769 | -105.1748
10.7806 | -105.1793
10.7842 | -105.1848
10.7879 | -105.1913
10.7915 | -105.1929
10.7952 | -105.1913
10.7988 | -105.1909
10.8025 | -105.1883
10.8062 | -105.1807
10.8098 | -105.1703
10.8135 | -105.1593
10.8171 | -105.1439
10.8208 | -105.1267
10.8245 | -105.114
10.8281 | -105.1013
10.8318 | -105.0889
10.8354 | -105.0813
10.8391 | -105.0687
10.8427 | -105.0598
10.8464 | -105.0532
10.8501 | -105.0504
10.8537 | -105.0534
10.8574 | -105.0596
10.861 | -105.0695
10.8647 | -105.0804
10.8684 | -105.0947
10.872 | -105.1101
10.8757 | -105.1245
10.8793 | -105.1398
10.883 | -105.1552
10.8866 | -105.1703
10.8903 | -105.1861
10.894 | -105.1982
10.8976 | -105.2078
10.9013 | -105.2186
10.9049 | -105.23
10.9086 | -105.2383
10.9123 | -105.2441
10.9159 | -105.2491
10.9196 | -105.2531
10.9232 | -105.2486
10.9269 | -105.2409
10.9305 | -105.233
10.9342 | -105.2267
10.9379 | -105.2211
10.9415 | -105.2091
10.9452 | -105.1933
10.9488 | -105.1802
10.9525 | -105.1605
10.9561 | -105.138
10.9598 | -105.1181
10.9635 | -105.1012
10.9671 | -105.0835
10.9708 | -105.0667
10.9744 | -105.0469
10.9781 | -105.0257
10.9818 | -105.0076
10.9854 | -104.9885
10.9891 | -104.9704
10.9927 | -104.9509
10.9964 | -104.9336
11 | -104.9137
11.0037 | -104.8946
11.0074 | -104.8753
11.011 | -104.8533
11.0147 | -104.8338
11.0183 | -104.8126
11.022 | -104.7922
11.0257 | -104.7754
11.0293 | -104.7622
11.033 | -104.7508
11.0366 | -104.7428
11.0403 | -104.7386
11.0439 | -104.74
11.0476 | -104.7475
11.0513 | -104.7577
11.0549 | -104.7731
11.0586 | -104.7897
11.0622 | -104.8101
11.0659 | -104.8328
11.0696 | -104.859
11.0732 | -104.8841
11.0769 | -104.9093
11.0805 | -104.9313
11.0842 | -104.955
11.0878 | -104.9786
11.0915 | -104.9946
11.0952 | -105.0069
11.0988 | -105.024
11.1025 | -105.0404
11.1061 | -105.0529
11.1098 | -105.063
11.1134 | -105.0752
11.1171 | -105.0875
11.1208 | -105.0937
11.1244 | -105.0986
11.1281 | -105.1051
11.1317 | -105.1123
11.1354 | -105.1223
11.1391 | -105.1309
11.1427 | -105.1397
11.1464 | -105.1508
11.15 | -105.1607
11.1537 | -105.1733
11.1573 | -105.1898
11.161 | -105.2075
11.1647 | -105.2207
11.1683 | -105.239
11.172 | -105.2611
11.1756 | -105.2799
11.1793 | -105.2984
11.183 | -105.3206
11.1866 | -105.3391
11.1903 | -105.3562
11.1939 | -105.374
11.1976 | -105.3927
11.2012 | -105.409
11.2049 | -105.4212
11.2086 | -105.4283
11.2122 | -105.4347
11.2159 | -105.4321
11.2195 | -105.4309
11.2232 | -105.4318
11.2269 | -105.4287
11.2305 | -105.4213
11.2342 | -105.4141
11.2378 | -105.4051
11.2415 | -105.392
11.2451 | -105.3798
11.2488 | -105.3626
11.2525 | -105.3412
11.2561 | -105.3217
11.2598 | -105.2993
11.2634 | -105.2716
11.2671 | -105.246
11.2707 | -105.2239
11.2744 | -105.2017
11.2781 | -105.1799
11.2817 | -105.1578
11.2854 | -105.1378
11.289 | -105.1155
11.2927 | -105.0974
11.2964 | -105.0788
11.3 | -105.0661
11.3037 | -105.0546
11.3073 | -105.0396
11.311 | -105.0296
11.3146 | -105.0233
11.3183 | -105.0191
11.322 | -105.0173
11.3256 | -105.021
11.3293 | -105.0287
11.3329 | -105.0329
11.3366 | -105.0427
11.3403 | -105.056
11.3439 | -105.0703
11.3476 | -105.0817
11.3512 | -105.0962
11.3549 | -105.1139
11.3585 | -105.1319
11.3622 | -105.1466
11.3659 | -105.161
11.3695 | -105.1791
11.3732 | -105.195
11.3768 | -105.2092
11.3805 | -105.2278
11.3842 | -105.2394
11.3878 | -105.2446
11.3915 | -105.2495
11.3951 | -105.2565
11.3988 | -105.2623
11.4024 | -105.2614
11.4061 | -105.2568
11.4098 | -105.2513
11.4134 | -105.2457
11.4171 | -105.241
11.4207 | -105.2363
11.4244 | -105.2291
11.428 | -105.2181
11.4317 | -105.2085
11.4354 | -105.2006
11.439 | -105.1921
11.4427 | -105.1834
11.4463 | -105.1749
11.45 | -105.1706
11.4537 | -105.1681
11.4573 | -105.1619
11.461 | -105.1594
11.4646 | -105.1527
11.4683 | -105.1482
11.4719 | -105.1448
11.4756 | -105.1361
11.4793 | -105.1282
11.4829 | -105.1181
11.4866 | -105.1086
11.4902 | -105.0974
11.4939 | -105.0871
11.4976 | -105.0755
11.5012 | -105.0602
11.5049 | -105.0434
11.5085 | -105.0303
11.5122 | -105.0169
11.5158 | -104.9994
11.5195 | -104.9831
11.5232 | -104.9663
11.5268 | -104.949
11.5305 | -104.9342
11.5341 | -104.9209
11.5378 | -104.9093
11.5415 | -104.8987
11.5451 | -104.8885
11.5488 | -104.882
11.5524 | -104.8783
11.5561 | -104.8773
11.5597 | -104.8799
11.5634 | -104.8848
11.5671 | -104.8961
11.5707 | -104.9098
11.5744 | -104.9262
11.578 | -104.9438
11.5817 | -104.963
11.5853 | -104.9836
11.589 | -105.0075
11.5927 | -105.0311
11.5963 | -105.0541
11.6 | -105.0755
11.6036 | -105.0993
11.6073 | -105.125
11.611 | -105.1521
11.6146 | -105.1815
11.6183 | -105.2073
11.6219 | -105.2327
11.6256 | -105.2589
11.6292 | -105.2816
11.6329 | -105.3043
11.6366 | -105.3283
11.6402 | -105.3513
11.6439 | -105.3771
11.6475 | -105.3968
11.6512 | -105.4154
11.6549 | -105.4337
11.6585 | -105.4529
11.6622 | -105.4705
11.6658 | -105.4916
11.6695 | -105.5108
11.6731 | -105.5307
11.6768 | -105.5441
11.6805 | -105.5571
11.6841 | -105.5669
11.6878 | -105.5703
11.6914 | -105.5786
11.6951 | -105.588
11.6988 | -105.5917
11.7024 | -105.5874
11.7061 | -105.585
11.7097 | -105.5853
11.7134 | -105.5841
11.717 | -105.5791
11.7207 | -105.572
11.7244 | -105.5634
11.728 | -105.5528
11.7317 | -105.5453
11.7353 | -105.5392
11.739 | -105.5307
11.7426 | -105.5187
11.7463 | -105.5074
11.75 | -105.4951
11.7536 | -105.4814
11.7573 | -105.468
11.7609 | -105.4601
11.7646 | -105.4502
11.7683 | -105.4353
11.7719 | -105.421
11.7756 | -105.4066
11.7792 | -105.3911
11.7829 | -105.3741
11.7865 | -105.3548
11.7902 | -105.3393
11.7939 | -105.3252
11.7975 | -105.3122
11.8012 | -105.3018
11.8048 | -105.2929
11.8085 | -105.2792
11.8122 | -105.2654
11.8158 | -105.2502
11.8195 | -105.2404
11.8231 | -105.2332
11.8268 | -105.2262
11.8304 | -105.2176
11.8341 | -105.2103
11.8378 | -105.2021
11.8414 | -105.1953
11.8451 | -105.1876
11.8487 | -105.1845
11.8524 | -105.1823
11.8561 | -105.1821
11.8597 | -105.1842
11.8634 | -105.1814
11.867 | -105.1772
11.8707 | -105.1741
11.8743 | -105.1751
11.878 | -105.1751
11.8817 | -105.1721
11.8853 | -105.1691
11.889 | -105.1685
11.8926 | -105.1691
11.8963 | -105.1731
11.9 | -105.1756
11.9036 | -105.1743
11.9073 | -105.1751
11.9109 | -105.1776
11.9146 | -105.1827
11.9182 | -105.1921
11.9219 | -105.1995
11.9256 | -105.2063
11.9292 | -105.2135
11.9329 | -105.2197
11.9365 | -105.2282
11.9402 | -105.2377
11.9438 | -105.246
11.9475 | -105.2528
11.9512 | -105.2597
11.9548 | -105.2673
11.9585 | -105.2778
11.9621 | -105.284
11.9658 | -105.2896
11.9695 | -105.294
11.9731 | -105.296
11.9768 | -105.2983
11.9804 | -105.3059
11.9841 | -105.3166
11.9877 | -105.3281
11.9914 | -105.3379
11.9951 | -105.3456
11.9987 | -105.3551
12.0024 | -105.3698
12.006 | -105.3847
12.0097 | -105.4023
12.0134 | -105.4202
12.017 | -105.4385
12.0207 | -105.4565
12.0243 | -105.4745
12.028 | -105.4948
12.0316 | -105.5133
12.0353 | -105.5328
12.039 | -105.5569
12.0426 | -105.58
12.0463 | -105.6003
12.0499 | -105.6183
12.0536 | -105.633
12.0572 | -105.6462
12.0609 | -105.6599
12.0646 | -105.6702
12.0682 | -105.6724
12.0719 | -105.6751
12.0755 | -105.6792
12.0792 | -105.6778
12.0829 | -105.6732
12.0865 | -105.6617
12.0902 | -105.6429
12.0938 | -105.6237
12.0975 | -105.6051
12.1011 | -105.5822
12.1048 | -105.5591
12.1085 | -105.5378
12.1121 | -105.5168
12.1158 | -105.4938
12.1194 | -105.4738
12.1231 | -105.456
12.1268 | -105.4421
12.1304 | -105.43
12.1341 | -105.4199
12.1377 | -105.4201
12.1414 | -105.4148
12.145 | -105.4105
12.1487 | -105.4094
12.1524 | -105.4099
12.156 | -105.409
12.1597 | -105.4094
12.1633 | -105.4089
12.167 | -105.4105
12.1707 | -105.416
12.1743 | -105.4207
12.178 | -105.4216
12.1816 | -105.4242
12.1853 | -105.4285
12.1889 | -105.4345
12.1926 | -105.4407
12.1963 | -105.4458
12.1999 | -105.4515
12.2036 | -105.4563
12.2072 | -105.4579
12.2109 | -105.462
12.2145 | -105.4641
12.2182 | -105.4669
12.2219 | -105.4721
12.2255 | -105.4787
12.2292 | -105.4861
12.2328 | -105.492
12.2365 | -105.4939
12.2402 | -105.4979
12.2438 | -105.5037
12.2475 | -105.507
12.2511 | -105.5122
12.2548 | -105.5155
12.2584 | -105.5193
12.2621 | -105.5243
12.2658 | -105.5322
12.2694 | -105.5406
12.2731 | -105.5454
12.2767 | -105.5496
12.2804 | -105.5504
12.2841 | -105.5482
12.2877 | -105.5516
12.2914 | -105.5547
12.295 | -105.5527
12.2987 | -105.5506
12.3023 | -105.5444
12.306 | -105.5376
12.3097 | -105.5326
12.3133 | -105.5244
12.317 | -105.5117
12.3206 | -105.4993
12.3243 | -105.4861
12.328 | -105.4687
12.3316 | -105.4491
12.3353 | -105.4338
12.3389 | -105.4167
12.3426 | -105.3995
12.3462 | -105.3811
12.3499 | -105.3617
12.3536 | -105.3407
12.3572 | -105.3191
12.3609 | -105.2981
12.3645 | -105.2721
12.3682 | -105.2472
12.3719 | -105.2242
12.3755 | -105.2011
12.3792 | -105.1789
12.3828 | -105.1585
12.3865 | -105.1398
12.3901 | -105.1197
12.3938 | -105.0981
12.3975 | -105.0794
12.4011 | -105.062
12.4048 | -105.0466
12.4084 | -105.0348
12.4121 | -105.0264
12.4157 | -105.0208
12.4194 | -105.0143
12.4231 | -105.012
12.4267 | -105.012
12.4304 | -105.0122
12.434 | -105.0194
12.4377 | -105.0296
12.4414 | -105.0392
12.445 | -105.0492
12.4487 | -105.0603
12.4523 | -105.0784
12.456 | -105.0986
12.4596 | -105.1187
12.4633 | -105.1379
12.467 | -105.1556
12.4706 | -105.1755
12.4743 | -105.1921
12.4779 | -105.209
12.4816 | -105.23
12.4853 | -105.248
12.4889 | -105.2674
12.4926 | -105.287
12.4962 | -105.3076
12.4999 | -105.3245
12.5035 | -105.3403
12.5072 | -105.3592
12.5109 | -105.3764
12.5145 | -105.3887
12.5182 | -105.3984
12.5218 | -105.4107
12.5255 | -105.4241
12.5292 | -105.4312
12.5328 | -105.4377
12.5365 | -105.4483
12.5401 | -105.4542
12.5438 | -105.4601
12.5474 | -105.462
12.5511 | -105.4659
12.5548 | -105.4692
12.5584 | -105.4693
12.5621 | -105.4668
12.5657 | -105.4646
12.5694 | -105.4615
12.573 | -105.4591
12.5767 | -105.4526
12.5804 | -105.4434
12.584 | -105.4335
12.5877 | -105.4246
12.5913 | -105.4139
12.595 | -105.3996
12.5987 | -105.3892
12.6023 | -105.377
12.606 | -105.363
12.6096 | -105.3483
12.6133 | -105.3317
12.6169 | -105.3154
12.6206 | -105.3033
12.6243 | -105.2894
12.6279 | -105.2713
12.6316 | -105.2509
12.6352 | -105.2308
12.6389 | -105.213
12.6426 | -105.1969
12.6462 | -105.1794
12.6499 | -105.1599
12.6535 | -105.1439
12.6572 | -105.1255
12.6608 | -105.1026
12.6645 | -105.0812
12.6682 | -105.064
12.6718 | -105.0448
12.6755 | -105.0288
12.6791 | -105.0116
12.6828 | -104.9971
12.6865 | -104.9847
12.6901 | -104.9724
12.6938 | -104.9627
12.6974 | -104.955
12.7011 | -104.9539
12.7047 | -104.9524
12.7084 | -104.9516
12.7121 | -104.9576
12.7157 | -104.9655
12.7194 | -104.9747
12.723 | -104.9865
12.7267 | -104.9998
12.7303 | -105.0157
12.734 | -105.033
12.7377 | -105.0516
12.7413 | -105.0733
12.745 | -105.0968
12.7486 | -105.121
12.7523 | -105.1434
12.756 | -105.1686
12.7596 | -105.1927
12.7633 | -105.2156
12.7669 | -105.2376
12.7706 | -105.2545
12.7742 | -105.2733
12.7779 | -105.2955
12.7816 | -105.3124
12.7852 | -105.3283
12.7889 | -105.3453
12.7925 | -105.3603
12.7962 | -105.3718
12.7999 | -105.3811
12.8035 | -105.3875
12.8072 | -105.3943
12.8108 | -105.4064
12.8145 | -105.4149
12.8181 | -105.4205
12.8218 | -105.4256
12.8255 | -105.4268
12.8291 | -105.4249
12.8328 | -105.4254
12.8364 | -105.4242
12.8401 | -105.4187
12.8438 | -105.4117
12.8474 | -105.4061
12.8511 | -105.3981
12.8547 | -105.3889
12.8584 | -105.3799
12.862 | -105.3673
12.8657 | -105.3556
12.8694 | -105.3438
12.873 | -105.3285
12.8767 | -105.3135
12.8803 | -105.296
12.884 | -105.2763
12.8876 | -105.2569
12.8913 | -105.2417
12.895 | -105.2264
12.8986 | -105.2084
12.9023 | -105.1908
12.9059 | -105.1763
12.9096 | -105.1633
12.9133 | -105.1505
12.9169 | -105.1343
12.9206 | -105.1174
12.9242 | -105.1063
12.9279 | -105.0952
12.9315 | -105.0841
12.9352 | -105.0733
12.9389 | -105.0666
12.9425 | -105.0572
12.9462 | -105.049
12.9498 | -105.0427
12.9535 | -105.0377
12.9572 | -105.0344
12.9608 | -105.0291
12.9645 | -105.023
12.9681 | -105.0207
12.9718 | -105.0184
12.9754 | -105.0135
12.9791 | -105.0106
12.9828 | -105.0126
12.9864 | -105.0145
12.9901 | -105.0147
12.9937 | -105.0142
12.9974 | -105.0123
13.0011 | -105.0143
13.0047 | -105.0197
13.0084 | -105.0228
13.012 | -105.0267
13.0157 | -105.0324
13.0193 | -105.0392
13.023 | -105.0487
13.0267 | -105.0623
13.0303 | -105.0746
13.034 | -105.0831
13.0376 | -105.093
13.0413 | -105.104
13.0449 | -105.1171
13.0486 | -105.1309
13.0523 | -105.1451
13.0559 | -105.1612
13.0596 | -105.1816
13.0632 | -105.2005
13.0669 | -105.2197
13.0706 | -105.243
13.0742 | -105.2633
13.0779 | -105.2796
13.0815 | -105.2968
13.0852 | -105.3139
13.0888 | -105.3295
13.0925 | -105.3442
13.0962 | -105.3593
13.0998 | -105.3745
13.1035 | -105.3893
13.1071 | -105.3982
13.1108 | -105.4017
13.1145 | -105.4047
13.1181 | -105.4058
13.1218 | -105.404
13.1254 | -105.3998
13.1291 | -105.3954
13.1327 | -105.3917
13.1364 | -105.3868
13.1401 | -105.3793
13.1437 | -105.3697
13.1474 | -105.3616
13.151 | -105.3546
13.1547 | -105.3465
13.1584 | -105.3404
13.162 | -105.3347
13.1657 | -105.3303
13.1693 | -105.3207
13.173 | -105.3167
13.1766 | -105.316
13.1803 | -105.3157
13.184 | -105.3156
13.1876 | -105.3139
13.1913 | -105.318
13.1949 | -105.3134
13.1986 | -105.3134
13.2022 | -105.3165
13.2059 | -105.3186
13.2096 | -105.3194
13.2132 | -105.3216
13.2169 | -105.324
13.2205 | -105.3277
13.2242 | -105.3275
13.2279 | -105.3308
13.2315 | -105.3343
13.2352 | -105.3354
13.2388 | -105.3369
13.2425 | -105.3373
13.2461 | -105.3352
13.2498 | -105.3321
13.2535 | -105.3338
13.2571 | -105.3353
13.2608 | -105.3299
13.2644 | -105.3218
13.2681 | -105.3172
13.2718 | -105.3131
13.2754 | -105.3092
13.2791 | -105.3031
13.2827 | -105.2908
13.2864 | -105.2744
13.29 | -105.2563
13.2937 | -105.2413
13.2974 | -105.2255
13.301 | -105.2089
13.3047 | -105.1899
13.3083 | -105.1687
13.312 | -105.1464
13.3157 | -105.1218
13.3193 | -105.0957
13.323 | -105.0669
13.3266 | -105.0386
13.3303 | -105.0119
13.3339 | -104.9893
13.3376 | -104.9653
13.3413 | -104.9402
13.3449 | -104.9173
13.3486 | -104.8982
13.3522 | -104.8826
13.3559 | -104.8665
13.3596 | -104.8571
13.3632 | -104.8471
13.3669 | -104.8379
13.3705 | -104.8322
13.3742 | -104.8303
13.3778 | -104.8317
13.3815 | -104.8315
13.3852 | -104.8362
13.3888 | -104.8476
13.3925 | -104.8582
13.3961 | -104.8699
13.3998 | -104.8829
13.4034 | -104.8958
13.4071 | -104.9103
13.4108 | -104.9271
13.4144 | -104.9459
13.4181 | -104.9697
13.4217 | -104.9917
13.4254 | -105.0143
13.4291 | -105.0393
13.4327 | -105.0646
13.4364 | -105.087
13.44 | -105.1103
13.4437 | -105.1325
13.4473 | -105.154
13.451 | -105.1773
13.4547 | -105.202
13.4583 | -105.2252
13.462 | -105.2404
13.4656 | -105.253
13.4693 | -105.2675
13.473 | -105.2801
13.4766 | -105.292
13.4803 | -105.3052
13.4839 | -105.3161
13.4876 | -105.3227
13.4912 | -105.3267
13.4949 | -105.3301
13.4986 | -105.3322
13.5022 | -105.3321
13.5059 | -105.3298
13.5095 | -105.3252
13.5132 | -105.3174
13.5168 | -105.3092
13.5205 | -105.2999
13.5242 | -105.2924
13.5278 | -105.2823
13.5315 | -105.2693
13.5351 | -105.2557
13.5388 | -105.2442
13.5425 | -105.2323
13.5461 | -105.2218
13.5498 | -105.2134
13.5534 | -105.2004
13.5571 | -105.1891
13.5607 | -105.1801
13.5644 | -105.1737
13.5681 | -105.1667
13.5717 | -105.1604
13.5754 | -105.1594
13.579 | -105.1591
13.5827 | -105.1583
13.5864 | -105.1593
13.59 | -105.1585
13.5937 | -105.1608
13.5973 | -105.1617
13.601 | -105.1575
13.6046 | -105.1509
13.6083 | -105.1505
13.612 | -105.1481
13.6156 | -105.1437
13.6193 | -105.1382
13.6229 | -105.1329
13.6266 | -105.1271
13.6303 | -105.122
13.6339 | -105.1144
13.6376 | -105.1046
13.6412 | -105.0894
13.6449 | -105.0739
13.6485 | -105.0611
13.6522 | -105.0475
13.6559 | -105.0347
13.6595 | -105.018
13.6632 | -104.9972
13.6668 | -104.9808
13.6705 | -104.967
13.6742 | -104.9537
13.6778 | -104.9413
13.6815 | -104.9309
13.6851 | -104.9247
13.6888 | -104.9184
13.6924 | -104.9179
13.6961 | -104.9219
13.6998 | -104.9305
13.7034 | -104.9437
13.7071 | -104.9611
13.7107 | -104.9809
13.7144 | -105.0035
13.718 | -105.0377
13.7217 | -105.0586
13.7254 | -105.0853
13.729 | -105.1169
13.7327 | -105.1493
13.7363 | -105.1788
13.74 | -105.2083
13.7437 | -105.2336
13.7473 | -105.2599
13.751 | -105.2855
13.7546 | -105.3088
13.7583 | -105.3304
13.7619 | -105.3501
13.7656 | -105.3666
13.7693 | -105.3826
13.7729 | -105.3934
13.7766 | -105.4011
13.7802 | -105.4057
13.7839 | -105.4055
13.7876 | -105.4044
13.7912 | -105.4024
13.7949 | -105.4002
13.7985 | -105.3975
13.8022 | -105.3938
13.8058 | -105.3816
13.8095 | -105.367
13.8132 | -105.3494
13.8168 | -105.3325
13.8205 | -105.3162
13.8241 | -105.2988
13.8278 | -105.2824
13.8315 | -105.2662
13.8351 | -105.2528
13.8388 | -105.2424
13.8424 | -105.2279
13.8461 | -105.2131
13.8497 | -105.2042
13.8534 | -105.1955
13.8571 | -105.19
13.8607 | -105.1882
13.8644 | -105.1853
13.868 | -105.1814
13.8717 | -105.1803
13.8753 | -105.1834
13.879 | -105.1884
13.8827 | -105.1969
13.8863 | -105.2064
13.89 | -105.2133
13.8936 | -105.2238
13.8973 | -105.234
13.901 | -105.2441
13.9046 | -105.2557
13.9083 | -105.2693
13.9119 | -105.2823
13.9156 | -105.2955
13.9192 | -105.3098
13.9229 | -105.322
13.9266 | -105.332
13.9302 | -105.3427
13.9339 | -105.3542
13.9375 | -105.365
13.9412 | -105.3737
13.9449 | -105.3872
13.9485 | -105.3979
13.9522 | -105.4081
13.9558 | -105.415
13.9595 | -105.408
13.9631 | -105.3943
13.9668 | -105.3866
13.9705 | -105.3819
13.9741 | -105.3762
13.9778 | -105.3684
13.9814 | -105.3606
13.9851 | -105.3532
13.9888 | -105.3463
13.9924 | -105.3342
13.9961 | -105.3198
13.9997 | -105.301
14.0034 | -105.2785
14.007 | -105.2572
14.0107 | -105.2334
14.0144 | -105.2092
14.018 | -105.1835
14.0217 | -105.1563
14.0253 | -105.1295
14.029 | -105.1016
14.0326 | -105.0736
14.0363 | -105.0456
14.04 | -105.0206
14.0436 | -105.0008
14.0473 | -104.9837
14.0509 | -104.9678
14.0546 | -104.9554
14.0583 | -104.9469
14.0619 | -104.9409
14.0656 | -104.9397
14.0692 | -104.943
14.0729 | -104.9523
14.0765 | -104.9648
14.0802 | -104.9819
14.0839 | -105.0053
14.0875 | -105.0356
14.0912 | -105.0721
14.0948 | -105.1098
14.0985 | -105.1523
14.1022 | -105.1969
14.1058 | -105.2415
14.1095 | -105.2886
14.1131 | -105.3374
14.1168 | -105.3868
14.1204 | -105.4389
14.1241 | -105.49
14.1278 | -105.5385
14.1314 | -105.5843
14.1351 | -105.6299
14.1387 | -105.6726
14.1424 | -105.7135
14.1461 | -105.751
14.1497 | -105.7858
14.1534 | -105.8193
14.157 | -105.8506
14.1607 | -105.8884
14.1643 | -105.9242
14.168 | -105.9445
14.1717 | -105.9547
14.1753 | -105.9652
14.179 | -105.9705
14.1826 | -105.9716
14.1863 | -105.9749
14.1899 | -105.9734
14.1936 | -105.966
14.1973 | -105.9568
14.2009 | -105.9522
14.2046 | -105.9468
14.2082 | -105.9373
14.2119 | -105.929
14.2156 | -105.919
14.2192 | -105.9069
14.2229 | -105.8935
14.2265 | -105.8771
14.2302 | -105.859
14.2338 | -105.8407
14.2375 | -105.8238
14.2412 | -105.8059
14.2448 | -105.7879
14.2485 | -105.7725
14.2521 | -105.7486
14.2558 | -105.7264
14.2595 | -105.7064
14.2631 | -105.6873
14.2668 | -105.6683
14.2704 | -105.6473
14.2741 | -105.6247
14.2777 | -105.6005
14.2814 | -105.578
14.2851 | -105.5534
14.2887 | -105.527
14.2924 | -105.4961
14.296 | -105.4663
14.2997 | -105.436
14.3034 | -105.4067
14.307 | -105.3793
14.3107 | -105.3539
14.3143 | -105.3257
14.318 | -105.2982
14.3216 | -105.2692
14.3253 | -105.2408
14.329 | -105.2133
14.3326 | -105.1863
14.3363 | -105.1627
14.3399 | -105.1357
14.3436 | -105.1124
14.3473 | -105.0903
14.3509 | -105.0702
14.3546 | -105.0531
14.3582 | -105.0376
14.3619 | -105.0259
14.3655 | -105.0182
14.3692 | -105.011
14.3729 | -105.0088
14.3765 | -105.007
14.3802 | -105.0045
14.3838 | -105.007
14.3875 | -105.0111
14.3911 | -105.014
14.3948 | -105.0198
14.3985 | -105.0274
14.4021 | -105.036
14.4058 | -105.0412
14.4094 | -105.047
14.4131 | -105.0538
14.4168 | -105.0594
14.4204 | -105.0644
14.4241 | -105.0693
14.4277 | -105.0775
14.4314 | -105.0876
14.435 | -105.0959
14.4387 | -105.1036
14.4424 | -105.1111
14.446 | -105.117
14.4497 | -105.1215
14.4533 | -105.126
14.457 | -105.1328
14.4607 | -105.1434
14.4643 | -105.1511
14.468 | -105.1604
14.4716 | -105.1678
14.4753 | -105.1755
14.4789 | -105.1832
14.4826 | -105.1924
14.4863 | -105.202
14.4899 | -105.2121
14.4936 | -105.2229
14.4972 | -105.2335
14.5009 | -105.2476
14.5046 | -105.2603
14.5082 | -105.2739
14.5119 | -105.2849
14.5155 | -105.2952
14.5192 | -105.306
14.5228 | -105.3156
14.5265 | -105.3251
14.5302 | -105.3368
14.5338 | -105.3471
14.5375 | -105.3589
14.5411 | -105.3714
14.5448 | -105.3824
14.5484 | -105.398
14.5521 | -105.4103
14.5558 | -105.4227
14.5594 | -105.4344
14.5631 | -105.4449
14.5667 | -105.4563
14.5704 | -105.4674
14.5741 | -105.4779
14.5777 | -105.4879
14.5814 | -105.4956
14.585 | -105.5071
14.5887 | -105.5173
14.5923 | -105.524
14.596 | -105.5302
14.5997 | -105.5335
14.6033 | -105.5355
14.607 | -105.5399
14.6106 | -105.5415
14.6143 | -105.5427
14.618 | -105.544
14.6216 | -105.5424
14.6253 | -105.5409
14.6289 | -105.5389
14.6326 | -105.535
14.6362 | -105.5306
14.6399 | -105.5251
14.6436 | -105.5231
14.6472 | -105.5224
14.6509 | -105.5188
14.6545 | -105.5166
14.6582 | -105.5158
14.6618 | -105.513
14.6655 | -105.5086
14.6692 | -105.5044
14.6728 | -105.5011
14.6765 | -105.4985
14.6801 | -105.4968
14.6838 | -105.4949
14.6875 | -105.497
14.6911 | -105.5003
14.6948 | -105.5017
14.6984 | -105.5039
14.7021 | -105.5038
14.7057 | -105.5019
14.7094 | -105.5002
14.7131 | -105.4973
14.7167 | -105.4937
14.7204 | -105.4896
14.724 | -105.4843
14.7277 | -105.4783
14.7314 | -105.4733
14.735 | -105.4672
14.7387 | -105.4582
14.7423 | -105.4466
14.746 | -105.4325
14.7496 | -105.4162
14.7533 | -105.3989
14.757 | -105.3798
14.7606 | -105.3592
14.7643 | -105.3371
14.7679 | -105.3165
14.7716 | -105.2943
14.7753 | -105.2712
14.7789 | -105.25
14.7826 | -105.2267
14.7862 | -105.2052
14.7899 | -105.187
14.7935 | -105.1667
14.7972 | -105.1443
14.8009 | -105.1266
14.8045 | -105.1125
14.8082 | -105.0969
14.8118 | -105.0838
14.8155 | -105.07
14.8192 | -105.0552
14.8228 | -105.0425
14.8265 | -105.0334
14.8301 | -105.0243
14.8338 | -105.019
14.8374 | -105.0147
14.8411 | -105.012
14.8448 | -105.0104
14.8484 | -105.0096
14.8521 | -105.0133
14.8557 | -105.0152
14.8594 | -105.0153
14.863 | -105.0206
14.8667 | -105.026
14.8704 | -105.0313
14.874 | -105.0341
14.8777 | -105.0358
14.8813 | -105.0376
14.885 | -105.042
14.8887 | -105.0458
14.8923 | -105.0516
14.896 | -105.0571
14.8996 | -105.0602
14.9033 | -105.0633
14.9069 | -105.0646
14.9106 | -105.0648
14.9143 | -105.0666
14.9179 | -105.0696
14.9216 | -105.0721
14.9252 | -105.0736
14.9289 | -105.0727
14.9326 | -105.0721
14.9362 | -105.0716
14.9399 | -105.0701
14.9435 | -105.0666
14.9472 | -105.0609
14.9508 | -105.0566
14.9545 | -105.0562
14.9582 | -105.0559
14.9618 | -105.0546
14.9655 | -105.0547
14.9691 | -105.0548
14.9728 | -105.0595
14.9765 | -105.0636
14.9801 | -105.0646
14.9838 | -105.0626
14.9874 | -105.0554
14.9911 | -105.0495
14.9947 | -105.0471
14.9984 | -105.047
15.0021 | -105.0436
15.0057 | -105.0404
15.0094 | -105.0404
15.013 | -105.0369
15.0167 | -105.0305
15.0203 | -105.0216
15.024 | -105.0138
15.0277 | -105.0049
15.0313 | -104.9961
15.035 | -104.9855
15.0386 | -104.9737
15.0423 | -104.959
15.046 | -104.9412
15.0496 | -104.924
15.0533 | -104.9084
15.0569 | -104.8932
15.0606 | -104.878
15.0642 | -104.8611
15.0679 | -104.8463
15.0716 | -104.8344
15.0752 | -104.8239
15.0789 | -104.8148
15.0825 | -104.8088
15.0862 | -104.8055
15.0899 | -104.8078
15.0935 | -104.8141
15.0972 | -104.8213
15.1008 | -104.8272
15.1045 | -104.8355
15.1081 | -104.844
15.1118 | -104.8581
15.1155 | -104.8772
15.1191 | -104.8974
15.1228 | -104.9209
15.1264 | -104.9437
15.1301 | -104.9716
15.1338 | -105.0027
15.1374 | -105.0333
15.1411 | -105.0665
15.1447 | -105.1011
15.1484 | -105.1343
15.152 | -105.1674
15.1557 | -105.1984
15.1594 | -105.2289
15.163 | -105.261
15.1667 | -105.2906
15.1703 | -105.317
15.174 | -105.3419
15.1776 | -105.3689
15.1813 | -105.3923
15.185 | -105.4145
15.1886 | -105.4332
15.1923 | -105.4517
15.1959 | -105.4694
15.1996 | -105.4884
15.2033 | -105.5106
15.2069 | -105.5244
15.2106 | -105.5295
15.2142 | -105.534
15.2179 | -105.5417
15.2215 | -105.5447
15.2252 | -105.5472
15.2289 | -105.552
15.2325 | -105.556
15.2362 | -105.5584
15.2398 | -105.5568
15.2435 | -105.5525
15.2472 | -105.5504
15.2508 | -105.5501
15.2545 | -105.5471
15.2581 | -105.5451
15.2618 | -105.5439
15.2654 | -105.5399
15.2691 | -105.5338
15.2728 | -105.5259
15.2764 | -105.5179
15.2801 | -105.5107
15.2837 | -105.4971
15.2874 | -105.4889
15.2911 | -105.4819
15.2947 | -105.4709
15.2984 | -105.4599
15.302 | -105.4471
15.3057 | -105.4352
15.3093 | -105.4239
15.313 | -105.4103
15.3167 | -105.3964
15.3203 | -105.382
15.324 | -105.3669
15.3276 | -105.353
15.3313 | -105.3353
15.3349 | -105.319
15.3386 | -105.3053
15.3423 | -105.2929
15.3459 | -105.2789
15.3496 | -105.2639
15.3532 | -105.2468
15.3569 | -105.2318
15.3606 | -105.2148
15.3642 | -105.1998
15.3679 | -105.1864
15.3715 | -105.1717
15.3752 | -105.1591
15.3788 | -105.1478
15.3825 | -105.1333
15.3862 | -105.1223
15.3898 | -105.11
15.3935 | -105.0982
15.3971 | -105.0877
15.4008 | -105.0747
15.4045 | -105.0618
15.4081 | -105.0524
15.4118 | -105.0449
15.4154 | -105.0354
15.4191 | -105.0266
15.4227 | -105.0202
15.4264 | -105.0145
15.4301 | -105.0104
15.4337 | -105.0048
15.4374 | -104.9981
15.441 | -104.9908
15.4447 | -104.985
15.4484 | -104.9767
15.452 | -104.9692
15.4557 | -104.9649
15.4593 | -104.9634
15.463 | -104.9632
15.4666 | -104.9604
15.4703 | -104.959
15.474 | -104.958
15.4776 | -104.9581
15.4813 | -104.9577
15.4849 | -104.9591
15.4886 | -104.9614
15.4922 | -104.9657
15.4959 | -104.97
15.4996 | -104.9729
15.5032 | -104.9782
15.5069 | -104.9805
15.5105 | -104.9837
15.5142 | -104.9874
15.5179 | -104.9924
15.5215 | -104.9981
15.5252 | -105.0073
15.5288 | -105.0173
15.5325 | -105.024
15.5361 | -105.0303
15.5398 | -105.0381
15.5435 | -105.0459
15.5471 | -105.0523
15.5508 | -105.0576
15.5544 | -105.0645
15.5581 | -105.0708
15.5618 | -105.0749
15.5654 | -105.0782
15.5691 | -105.082
15.5727 | -105.0858
15.5764 | -105.0907
15.58 | -105.0947
15.5837 | -105.0995
15.5874 | -105.101
15.591 | -105.101
15.5947 | -105.1026
15.5983 | -105.1025
15.602 | -105.1037
15.6057 | -105.1067
15.6093 | -105.1064
15.613 | -105.1051
15.6166 | -105.102
15.6203 | -105.0972
15.6239 | -105.0931
15.6276 | -105.0902
15.6313 | -105.0832
15.6349 | -105.077
15.6386 | -105.0732
15.6422 | -105.0673
15.6459 | -105.0578
15.6495 | -105.0485
15.6532 | -105.0424
15.6569 | -105.0359
15.6605 | -105.0267
15.6642 | -105.021
15.6678 | -105.0137
15.6715 | -105.0077
15.6752 | -104.9981
15.6788 | -104.9846
15.6825 | -104.9747
15.6861 | -104.9634
15.6898 | -104.949
15.6934 | -104.9357
15.6971 | -104.9261
15.7008 | -104.9145
15.7044 | -104.9044
15.7081 | -104.8937
15.7117 | -104.8824
15.7154 | -104.8722
15.7191 | -104.8633
15.7227 | -104.8554
15.7264 | -104.8462
15.73 | -104.837
15.7337 | -104.8283
15.7373 | -104.8206
15.741 | -104.8122
15.7447 | -104.8041
15.7483 | -104.7986
15.752 | -104.7939
15.7556 | -104.7936
15.7593 | -104.7942
15.763 | -104.7943
15.7666 | -104.7965
15.7703 | -104.8002
15.7739 | -104.8052
15.7776 | -104.8094
15.7812 | -104.812
15.7849 | -104.8176
15.7886 | -104.8232
15.7922 | -104.8305
15.7959 | -104.8343
15.7995 | -104.8379
15.8032 | -104.8444
15.8069 | -104.8502
15.8105 | -104.8543
15.8142 | -104.8581
15.8178 | -104.8641
15.8215 | -104.8708
15.8251 | -104.8734
15.8288 | -104.8717
15.8325 | -104.8666
15.8361 | -104.8674
15.8398 | -104.8686
15.8434 | -104.8703
15.8471 | -104.8717
15.8507 | -104.8752
15.8544 | -104.8771
15.8581 | -104.8781
15.8617 | -104.8755
15.8654 | -104.8747
15.869 | -104.8742
15.8727 | -104.8745
15.8764 | -104.8734
15.88 | -104.87
15.8837 | -104.8651
15.8873 | -104.8639
15.891 | -104.8608
15.8946 | -104.8544
15.8983 | -104.8489
15.902 | -104.8464
15.9056 | -104.8428
15.9093 | -104.8372
15.9129 | -104.8312
15.9166 | -104.8241
15.9203 | -104.8143
15.9239 | -104.806
15.9276 | -104.7983
15.9312 | -104.7904
15.9349 | -104.7824
15.9385 | -104.7745
15.9422 | -104.7659
15.9459 | -104.7569
15.9495 | -104.7462
15.9532 | -104.7376
15.9568 | -104.7272
15.9605 | -104.7189
15.9642 | -104.7093
15.9678 | -104.7008
15.9715 | -104.6904
15.9751 | -104.6805
15.9788 | -104.6719
15.9824 | -104.6597
15.9861 | -104.649
15.9898 | -104.6383
15.9934 | -104.6266
15.9971 | -104.6164
16.0007 | -104.6068
16.0044 | -104.5975
16.008 | -104.5925
16.0117 | -104.5866
16.0154 | -104.583
16.019 | -104.5811
16.0227 | -104.5821
16.0263 | -104.5808
16.03 | -104.5796
16.0337 | -104.5829
16.0373 | -104.59
16.041 | -104.5974
16.0446 | -104.6008
16.0483 | -104.6073
16.0519 | -104.6189
16.0556 | -104.6347
16.0593 | -104.6491
16.0629 | -104.6615
16.0666 | -104.6702
16.0702 | -104.6796
16.0739 | -104.6922
16.0776 | -104.7039
16.0812 | -104.7142
16.0849 | -104.7243
16.0885 | -104.7338
16.0922 | -104.7454
16.0958 | -104.7548
16.0995 | -104.7643
16.1032 | -104.7717
16.1068 | -104.779
16.1105 | -104.7848
16.1141 | -104.7904
16.1178 | -104.7915
16.1214 | -104.7917
16.1251 | -104.7958
16.1288 | -104.7983
16.1324 | -104.7987
16.1361 | -104.8041
16.1397 | -104.8057
16.1434 | -104.8086
16.1471 | -104.8129
16.1507 | -104.8148
16.1544 | -104.8127
16.158 | -104.8116
16.1617 | -104.8103
16.1653 | -104.8118
16.169 | -104.8131
16.1727 | -104.8135
16.1763 | -104.8161
16.18 | -104.8218
16.1836 | -104.8241
16.1873 | -104.8268
16.191 | -104.8285
16.1946 | -104.8309
16.1983 | -104.8335
16.2019 | -104.8406
16.2056 | -104.8443
16.2092 | -104.8482
16.2129 | -104.8558
16.2166 | -104.8643
16.2202 | -104.8753
16.2239 | -104.8849
16.2275 | -104.8947
16.2312 | -104.9045
16.2349 | -104.9176
16.2385 | -104.9263
16.2422 | -104.9367
16.2458 | -104.9493
16.2495 | -104.963
16.2531 | -104.9762
16.2568 | -104.993
16.2605 | -105.0088
16.2641 | -105.0227
16.2678 | -105.0377
16.2714 | -105.0551
16.2751 | -105.0743
16.2788 | -105.0873
16.2824 | -105.0996
16.2861 | -105.1151
16.2897 | -105.131
16.2934 | -105.1457
16.297 | -105.1619
16.3007 | -105.1765
16.3044 | -105.1865
16.308 | -105.1939
16.3117 | -105.202
16.3153 | -105.2105
16.319 | -105.2198
16.3226 | -105.2271
16.3263 | -105.2327
16.33 | -105.2373
16.3336 | -105.2431
16.3373 | -105.2497
16.3409 | -105.2545
16.3446 | -105.2562
16.3483 | -105.261
16.3519 | -105.2664
16.3556 | -105.2688
16.3592 | -105.2705
16.3629 | -105.2707
16.3665 | -105.272
16.3702 | -105.27
16.3739 | -105.2703
16.3775 | -105.272
16.3812 | -105.2753
16.3848 | -105.2791
16.3885 | -105.2831
16.3922 | -105.2877
16.3958 | -105.2928
16.3995 | -105.2958
16.4031 | -105.3015
16.4068 | -105.3064
16.4104 | -105.3101
16.4141 | -105.3138
16.4178 | -105.3213
16.4214 | -105.3308
16.4251 | -105.3396
16.4287 | -105.3481
16.4324 | -105.3564
16.4361 | -105.3631
16.4397 | -105.3726
16.4434 | -105.3811
16.447 | -105.3911
16.4507 | -105.4018
16.4543 | -105.4085
16.458 | -105.4176
16.4617 | -105.4253
16.4653 | -105.4337
16.469 | -105.4373
16.4726 | -105.4422
16.4763 | -105.4449
16.4799 | -105.4465
16.4836 | -105.4474
16.4873 | -105.4496
16.4909 | -105.4513
16.4946 | -105.4518
16.4982 | -105.4493
16.5019 | -105.4432
16.5056 | -105.4342
16.5092 | -105.4244
16.5129 | -105.4155
16.5165 | -105.3993
16.5202 | -105.3803
16.5238 | -105.3612
16.5275 | -105.3438
16.5312 | -105.3252
16.5348 | -105.3067
16.5385 | -105.2859
16.5421 | -105.267
16.5458 | -105.2453
16.5495 | -105.2215
16.5531 | -105.1979
16.5568 | -105.1744
16.5604 | -105.1526
16.5641 | -105.1317
16.5677 | -105.1109
16.5714 | -105.0895
16.5751 | -105.0721
16.5787 | -105.0592
16.5824 | -105.0454
16.586 | -105.032
16.5897 | -105.0212
16.5934 | -105.0131
16.597 | -105.0081
16.6007 | -105.0072
16.6043 | -105.0064
16.608 | -105.0073
16.6116 | -105.0113
16.6153 | -105.0186
16.619 | -105.0275
16.6226 | -105.039
16.6263 | -105.0526
16.6299 | -105.0658
16.6336 | -105.0808
16.6372 | -105.0966
16.6409 | -105.1118
16.6446 | -105.1302
16.6482 | -105.1495
16.6519 | -105.17
16.6555 | -105.1887
16.6592 | -105.2038
16.6629 | -105.2202
16.6665 | -105.2385
16.6702 | -105.254
16.6738 | -105.2653
16.6775 | -105.2751
16.6811 | -105.2836
16.6848 | -105.2898
16.6885 | -105.2933
16.6921 | -105.2954
16.6958 | -105.2954
16.6994 | -105.2931
16.7031 | -105.2886
16.7068 | -105.2849
16.7104 | -105.2799
16.7141 | -105.2739
16.7177 | -105.2655
16.7214 | -105.2557
16.725 | -105.2448
16.7287 | -105.2341
16.7324 | -105.2217
16.736 | -105.2087
16.7397 | -105.1942
16.7433 | -105.1795
16.747 | -105.1654
16.7507 | -105.1547
16.7543 | -105.1461
16.758 | -105.1351
16.7616 | -105.1256
16.7653 | -105.1175
16.7689 | -105.1098
16.7726 | -105.1014
16.7763 | -105.0955
16.7799 | -105.0906
16.7836 | -105.0861
16.7872 | -105.0818
16.7909 | -105.0767
16.7945 | -105.0738
16.7982 | -105.0691
16.8019 | -105.0681
16.8055 | -105.067
16.8092 | -105.0679
16.8128 | -105.0698
16.8165 | -105.0704
16.8202 | -105.0745
16.8238 | -105.0807
16.8275 | -105.0859
16.8311 | -105.0888
16.8348 | -105.093
16.8384 | -105.0979
16.8421 | -105.1024
16.8458 | -105.108
16.8494 | -105.1126
16.8531 | -105.1162
16.8567 | -105.1209
16.8604 | -105.1262
16.8641 | -105.1321
16.8677 | -105.1404
16.8714 | -105.1459
16.875 | -105.1518
16.8787 | -105.1588
16.8823 | -105.1631
16.886 | -105.1719
16.8897 | -105.1771
16.8933 | -105.1809
16.897 | -105.1869
16.9006 | -105.1976
16.9043 | -105.2069
16.908 | -105.2146
16.9116 | -105.2247
16.9153 | -105.2378
16.9189 | -105.252
16.9226 | -105.2666
16.9262 | -105.2799
16.9299 | -105.2971
16.9336 | -105.3165
16.9372 | -105.3362
16.9409 | -105.3546
16.9445 | -105.3711
16.9482 | -105.3872
16.9519 | -105.4059
16.9555 | -105.4249
16.9592 | -105.4428
16.9628 | -105.4601
16.9665 | -105.4749
16.9701 | -105.4889
16.9738 | -105.5044
16.9775 | -105.5196
16.9811 | -105.5321
16.9848 | -105.5427
16.9884 | -105.5522
16.9921 | -105.5567
16.9957 | -105.5606
16.9994 | -105.5593
17.0031 | -105.5526
17.0067 | -105.5442
17.0104 | -105.5334
17.014 | -105.521
17.0177 | -105.5081
17.0214 | -105.4939
17.025 | -105.4786
17.0287 | -105.4612
17.0323 | -105.4425
17.036 | -105.4216
17.0396 | -105.402
17.0433 | -105.374
17.047 | -105.3474
17.0506 | -105.3191
17.0543 | -105.2865
17.0579 | -105.257
17.0616 | -105.2258
17.0653 | -105.1972
17.0689 | -105.1711
17.0726 | -105.1483
17.0762 | -105.1253
17.0799 | -105.1025
17.0835 | -105.0796
17.0872 | -105.0589
17.0909 | -105.0399
17.0945 | -105.0231
17.0982 | -105.0079
17.1018 | -104.9939
17.1055 | -104.981
17.1091 | -104.9667
17.1128 | -104.9561
17.1165 | -104.9472
17.1201 | -104.938
17.1238 | -104.9332
17.1274 | -104.9293
17.1311 | -104.9245
17.1348 | -104.9212
17.1384 | -104.9182
17.1421 | -104.9166
17.1457 | -104.915
17.1494 | -104.913
17.153 | -104.914
17.1567 | -104.9152
17.1604 | -104.9173
17.164 | -104.9177
17.1677 | -104.9218
17.1713 | -104.9231
17.175 | -104.9221
17.1787 | -104.9248
17.1823 | -104.9277
17.186 | -104.9336
17.1896 | -104.9422
17.1933 | -104.9508
17.1969 | -104.9558
17.2006 | -104.9621
17.2043 | -104.9683
17.2079 | -104.9738
17.2116 | -104.9817
17.2152 | -104.9892
17.2189 | -104.9989
17.2226 | -105.0109
17.2262 | -105.0253
17.2299 | -105.0388
17.2335 | -105.0503
17.2372 | -105.0626
17.2408 | -105.0758
17.2445 | -105.0909
17.2482 | -105.1097
17.2518 | -105.128
17.2555 | -105.1485
17.2591 | -105.1676
17.2628 | -105.1827
17.2664 | -105.1995
17.2701 | -105.2167
17.2738 | -105.2343
17.2774 | -105.2534
17.2811 | -105.2736
17.2847 | -105.2933
17.2884 | -105.3165
17.2921 | -105.3325
17.2957 | -105.3451
17.2994 | -105.3644
17.303 | -105.3813
17.3067 | -105.3926
17.3103 | -105.4005
17.314 | -105.4077
17.3177 | -105.4133
17.3213 | -105.4166
17.325 | -105.4176
17.3286 | -105.4191
17.3323 | -105.4199
17.336 | -105.4199
17.3396 | -105.4162
17.3433 | -105.4099
17.3469 | -105.4021
17.3506 | -105.3938
17.3542 | -105.3876
17.3579 | -105.3793
17.3616 | -105.3674
17.3652 | -105.3572
17.3689 | -105.344
17.3725 | -105.3289
17.3762 | -105.3163
17.3799 | -105.3042
17.3835 | -105.2925
17.3872 | -105.2793
17.3908 | -105.2664
17.3945 | -105.2559
17.3981 | -105.2437
17.4018 | -105.2323
17.4055 | -105.2221
17.4091 | -105.2175
17.4128 | -105.2277
17.4164 | -105.2249
17.4201 | -105.2267
17.4238 | -105.2277
17.4274 | -105.2291
17.4311 | -105.233
17.4347 | -105.2389
17.4384 | -105.2463
17.442 | -105.2551
17.4457 | -105.2653
17.4494 | -105.2746
17.453 | -105.2865
17.4567 | -105.301
17.4603 | -105.316
17.464 | -105.331
17.4676 | -105.3445
17.4713 | -105.355
17.475 | -105.3669
17.4786 | -105.38
17.4823 | -105.3926
17.4859 | -105.4055
17.4896 | -105.418
17.4933 | -105.4311
17.4969 | -105.4401
17.5006 | -105.445
17.5042 | -105.4473
17.5079 | -105.4541
17.5115 | -105.4621
17.5152 | -105.4651
17.5189 | -105.4688
17.5225 | -105.4718
17.5262 | -105.4692
17.5298 | -105.4616
17.5335 | -105.4534
17.5372 | -105.4448
17.5408 | -105.4416
17.5445 | -105.438
17.5481 | -105.4273
17.5518 | -105.4144
17.5554 | -105.4043
17.5591 | -105.3938
17.5628 | -105.3822
17.5664 | -105.3675
17.5701 | -105.3536
17.5737 | -105.3387
17.5774 | -105.3199
17.5811 | -105.2967
17.5847 | -105.2723
17.5884 | -105.2512
17.592 | -105.2304
17.5957 | -105.2067
17.5993 | -105.1814
17.603 | -105.1544
17.6067 | -105.128
17.6103 | -105.0991
17.614 | -105.0722
17.6176 | -105.0453
17.6213 | -105.0198
17.6249 | -104.9945
17.6286 | -104.969
17.6323 | -104.9442
17.6359 | -104.9208
17.6396 | -104.8985
17.6432 | -104.877
17.6469 | -104.8558
17.6506 | -104.8373
17.6542 | -104.8189
17.6579 | -104.8009
17.6615 | -104.7853
17.6652 | -104.7714
17.6688 | -104.7582
17.6725 | -104.747
17.6762 | -104.7404
17.6798 | -104.733
17.6835 | -104.7276
17.6871 | -104.7233
17.6908 | -104.7213
17.6945 | -104.7217
17.6981 | -104.7252
17.7018 | -104.7312
17.7054 | -104.7371
17.7091 | -104.7446
17.7127 | -104.7524
17.7164 | -104.7617
17.7201 | -104.771
17.7237 | -104.7808
17.7274 | -104.7926
17.731 | -104.8039
17.7347 | -104.8174
17.7384 | -104.8314
17.742 | -104.8444
17.7457 | -104.8569
17.7493 | -104.8696
17.753 | -104.8825
17.7566 | -104.8958
17.7603 | -104.9056
17.764 | -104.9161
17.7676 | -104.9265
17.7713 | -104.9351
17.7749 | -104.9423
17.7786 | -104.9517
17.7822 | -104.9585
17.7859 | -104.9641
17.7896 | -104.9676
17.7932 | -104.9687
17.7969 | -104.9688
17.8005 | -104.9689
17.8042 | -104.966
17.8079 | -104.962
17.8115 | -104.9567
17.8152 | -104.9531
17.8188 | -104.9491
17.8225 | -104.9413
17.8261 | -104.9341
17.8298 | -104.9307
17.8335 | -104.9276
17.8371 | -104.9239
17.8408 | -104.9129
17.8444 | -104.9007
17.8481 | -104.8926
17.8518 | -104.886
17.8554 | -104.8802
17.8591 | -104.8751
17.8627 | -104.8706
17.8664 | -104.8614
17.87 | -104.8525
17.8737 | -104.8441
17.8774 | -104.8368
17.881 | -104.8321
17.8847 | -104.829
17.8883 | -104.8237
17.892 | -104.8185
17.8957 | -104.8147
17.8993 | -104.8135
17.903 | -104.8108
17.9066 | -104.8062
17.9103 | -104.8007
17.9139 | -104.7948
17.9176 | -104.7912
17.9213 | -104.7894
17.9249 | -104.7884
17.9286 | -104.7861
17.9322 | -104.7857
17.9359 | -104.7846
17.9395 | -104.7828
17.9432 | -104.7851
17.9469 | -104.786
17.9505 | -104.7866
17.9542 | -104.7881
17.9578 | -104.791
17.9615 | -104.7928
17.9652 | -104.7945
17.9688 | -104.7977
17.9725 | -104.8007
17.9761 | -104.8037
17.9798 | -104.8085
17.9834 | -104.8116
17.9871 | -104.8141
17.9908 | -104.8166
17.9944 | -104.8198
17.9981 | -104.8236
18.0017 | -104.8287
18.0054 | -104.8345
18.0091 | -104.8411
18.0127 | -104.8465
18.0164 | -104.8532
18.02 | -104.859
18.0237 | -104.8659
18.0273 | -104.8717
18.031 | -104.8776
18.0347 | -104.8844
18.0383 | -104.8925
18.042 | -104.8993
18.0456 | -104.9035
18.0493 | -104.9054
18.053 | -104.9078
18.0566 | -104.9106
18.0603 | -104.9131
18.0639 | -104.9155
18.0676 | -104.9169
18.0712 | -104.9147
18.0749 | -104.9136
18.0786 | -104.9145
18.0822 | -104.9113
18.0859 | -104.9047
18.0895 | -104.8977
18.0932 | -104.891
18.0968 | -104.889
18.1005 | -104.8869
18.1042 | -104.8827
18.1078 | -104.8751
18.1115 | -104.8682
18.1151 | -104.8612
18.1188 | -104.855
18.1225 | -104.8492
18.1261 | -104.8436
18.1298 | -104.8383
18.1334 | -104.8311
18.1371 | -104.8239
18.1407 | -104.8155
18.1444 | -104.808
18.1481 | -104.8005
18.1517 | -104.7932
18.1554 | -104.7842
18.159 | -104.7745
18.1627 | -104.7679
18.1664 | -104.7599
18.17 | -104.7507
18.1737 | -104.7397
18.1773 | -104.7276
18.181 | -104.7167
18.1846 | -104.7063
18.1883 | -104.6935
18.192 | -104.6788
18.1956 | -104.6658
18.1993 | -104.6504
18.2029 | -104.6343
18.2066 | -104.6175
18.2103 | -104.6005
18.2139 | -104.5862
18.2176 | -104.5741
18.2212 | -104.565
18.2249 | -104.5557
18.2285 | -104.5448
18.2322 | -104.5357
18.2359 | -104.5299
18.2395 | -104.5254
18.2432 | -104.5212
18.2468 | -104.5173
18.2505 | -104.5106
18.2541 | -104.5012
18.2578 | -104.494
18.2615 | -104.4884
18.2651 | -104.4845
18.2688 | -104.4816
18.2724 | -104.479
18.2761 | -104.4758
18.2798 | -104.4733
18.2834 | -104.472
18.2871 | -104.4722
18.2907 | -104.4716
18.2944 | -104.4687
18.298 | -104.4666
18.3017 | -104.4641
18.3054 | -104.4644
18.309 | -104.4635
18.3127 | -104.4637
18.3163 | -104.4653
18.32 | -104.4674
18.3237 | -104.4681
18.3273 | -104.4681
18.331 | -104.4694
18.3346 | -104.4715
18.3383 | -104.4729
18.3419 | -104.4785
18.3456 | -104.4834
18.3493 | -104.4882
18.3529 | -104.4931
18.3566 | -104.4973
18.3602 | -104.5018
18.3639 | -104.5074
18.3676 | -104.5149
18.3712 | -104.5203
18.3749 | -104.5259
18.3785 | -104.534
18.3822 | -104.5434
18.3858 | -104.5525
18.3895 | -104.5587
18.3932 | -104.5685
18.3968 | -104.5804
18.4005 | -104.5937
18.4041 | -104.6083
18.4078 | -104.6221
18.4114 | -104.6376
18.4151 | -104.6546
18.4188 | -104.6724
18.4224 | -104.6907
18.4261 | -104.7113
18.4297 | -104.7319
18.4334 | -104.7547
18.4371 | -104.7799
18.4407 | -104.8014
18.4444 | -104.8228
18.448 | -104.8458
18.4517 | -104.8695
18.4553 | -104.8919
18.459 | -104.9142
18.4627 | -104.9357
18.4663 | -104.9567
18.47 | -104.9709
18.4736 | -104.9932
18.4773 | -105.0104
18.481 | -105.026
18.4846 | -105.0396
18.4883 | -105.0514
18.4919 | -105.0647
18.4956 | -105.0749
18.4992 | -105.0828
18.5029 | -105.0884
18.5066 | -105.0955
18.5102 | -105.1026
18.5139 | -105.1098
18.5175 | -105.1156
18.5212 | -105.1166
18.5249 | -105.1137
18.5285 | -105.1084
18.5322 | -105.1025
18.5358 | -105.0949
18.5395 | -105.0834
18.5431 | -105.0721
18.5468 | -105.0605
18.5505 | -105.0456
18.5541 | -105.0318
18.5578 | -105.0187
18.5614 | -105.0068
18.5651 | -104.9942
18.5688 | -104.9799
18.5724 | -104.968
18.5761 | -104.9526
18.5797 | -104.9351
18.5834 | -104.9184
18.587 | -104.9003
18.5907 | -104.8841
18.5944 | -104.8702
18.598 | -104.8567
18.6017 | -104.8444
18.6053 | -104.8291
18.609 | -104.8159
18.6126 | -104.8039
18.6163 | -104.7931
18.62 | -104.782
18.6236 | -104.7737
18.6273 | -104.7657
18.6309 | -104.7576
18.6346 | -104.7498
18.6383 | -104.7444
18.6419 | -104.7391
18.6456 | -104.7335
18.6492 | -104.7284
18.6529 | -104.7243
18.6565 | -104.719
18.6602 | -104.7137
18.6639 | -104.7098
18.6675 | -104.7073
18.6712 | -104.7067
18.6748 | -104.7047
18.6785 | -104.701
18.6822 | -104.6989
18.6858 | -104.6973
18.6895 | -104.693
18.6931 | -104.6887
18.6968 | -104.6833
18.7004 | -104.6778
18.7041 | -104.6759
18.7078 | -104.6747
18.7114 | -104.6726
18.7151 | -104.6696
18.7187 | -104.6679
18.7224 | -104.6666
18.7261 | -104.6653
18.7297 | -104.6629
18.7334 | -104.6586
18.737 | -104.6533
18.7407 | -104.6463
18.7443 | -104.6383
18.748 | -104.6294
18.7517 | -104.6219
18.7553 | -104.6152
18.759 | -104.6078
18.7626 | -104.5991
18.7663 | -104.589
18.7699 | -104.5789
18.7736 | -104.5652
18.7773 | -104.5528
18.7809 | -104.5402
18.7846 | -104.529
18.7882 | -104.5174
18.7919 | -104.5066
18.7956 | -104.4957
18.7992 | -104.4849
18.8029 | -104.4761
18.8065 | -104.4701
18.8102 | -104.4633
18.8138 | -104.4532
18.8175 | -104.4468
18.8212 | -104.4418
18.8248 | -104.438
18.8285 | -104.4334
18.8321 | -104.4306
18.8358 | -104.4273
18.8395 | -104.4246
18.8431 | -104.424
18.8468 | -104.4225
18.8504 | -104.4236
18.8541 | -104.4288
18.8577 | -104.4327
18.8614 | -104.4372
18.8651 | -104.442
18.8687 | -104.4466
18.8724 | -104.4535
18.876 | -104.4595
18.8797 | -104.4665
18.8834 | -104.4732
18.887 | -104.4807
18.8907 | -104.4886
18.8943 | -104.4969
18.898 | -104.5053
18.9016 | -104.5162
18.9053 | -104.5259
18.909 | -104.5363
18.9126 | -104.5479
18.9163 | -104.5582
18.9199 | -104.5667
18.9236 | -104.5757
18.9272 | -104.5818
18.9309 | -104.5882
18.9346 | -104.5976
18.9382 | -104.606
18.9419 | -104.6131
18.9455 | -104.6239
18.9492 | -104.6315
18.9529 | -104.6371
18.9565 | -104.6474
18.9602 | -104.657
18.9638 | -104.6646
18.9675 | -104.6696
18.9711 | -104.6755
18.9748 | -104.682
18.9785 | -104.6885
18.9821 | -104.6949
18.9858 | -104.6976
18.9894 | -104.7005
18.9931 | -104.699
18.9968 | -104.6946
19.0004 | -104.6952
19.0041 | -104.6935
19.0077 | -104.6908
19.0114 | -104.6881
19.015 | -104.684
19.0187 | -104.6792
19.0224 | -104.675
19.026 | -104.6701
19.0297 | -104.6645
19.0333 | -104.66
19.037 | -104.6554
19.0407 | -104.653
19.0443 | -104.6521
19.048 | -104.6493
19.0516 | -104.6485
19.0553 | -104.6501
19.0589 | -104.6518
19.0626 | -104.6529
19.0663 | -104.6544
19.0699 | -104.6584
19.0736 | -104.6638
19.0772 | -104.6702
19.0809 | -104.6764
19.0845 | -104.6872
19.0882 | -104.6986
19.0919 | -104.7102
19.0955 | -104.7231
19.0992 | -104.7362
19.1028 | -104.7522
19.1065 | -104.7709
19.1102 | -104.7889
19.1138 | -104.8077
19.1175 | -104.8279
19.1211 | -104.8487
19.1248 | -104.8665
19.1284 | -104.8834
19.1321 | -104.8997
19.1358 | -104.9187
19.1394 | -104.9378
19.1431 | -104.9558
19.1467 | -104.9728
19.1504 | -104.9881
19.1541 | -105.0029
19.1577 | -105.0203
19.1614 | -105.0352
19.165 | -105.0479
19.1687 | -105.0617
19.1723 | -105.0756
19.176 | -105.0877
19.1797 | -105.0996
19.1833 | -105.1123
19.187 | -105.1212
19.1906 | -105.1293
19.1943 | -105.1388
19.198 | -105.1488
19.2016 | -105.1559
19.2053 | -105.1592
19.2089 | -105.1619
19.2126 | -105.1655
19.2162 | -105.1652
19.2199 | -105.1617
19.2236 | -105.1594
19.2272 | -105.1538
19.2309 | -105.1481
19.2345 | -105.1453
19.2382 | -105.1423
19.2418 | -105.1347
19.2455 | -105.1232
19.2492 | -105.1107
19.2528 | -105.0972
19.2565 | -105.0883
19.2601 | -105.0773
19.2638 | -105.0672
19.2675 | -105.0561
19.2711 | -105.0428
19.2748 | -105.0309
19.2784 | -105.0219
19.2821 | -105.0124
19.2857 | -105.0008
19.2894 | -104.9909
19.2931 | -104.9804
19.2967 | -104.9646
19.3004 | -104.9486
19.304 | -104.9355
19.3077 | -104.9257
19.3114 | -104.9154
19.315 | -104.9034
19.3187 | -104.8927
19.3223 | -104.8829
19.326 | -104.8739
19.3296 | -104.8632
19.3333 | -104.8523
19.337 | -104.843
19.3406 | -104.8343
19.3443 | -104.8233
19.3479 | -104.8113
19.3516 | -104.801
19.3553 | -104.7929
19.3589 | -104.7827
19.3626 | -104.7713
19.3662 | -104.762
19.3699 | -104.752
19.3735 | -104.7441
19.3772 | -104.7356
19.3809 | -104.7261
19.3845 | -104.7186
19.3882 | -104.7108
19.3918 | -104.7059
19.3955 | -104.698
19.3991 | -104.6913
19.4028 | -104.6882
19.4065 | -104.6836
19.4101 | -104.6778
19.4138 | -104.6727
19.4174 | -104.6672
19.4211 | -104.6636
19.4248 | -104.6635
19.4284 | -104.665
19.4321 | -104.6657
19.4357 | -104.6677
19.4394 | -104.6695
19.443 | -104.6708
19.4467 | -104.6713
19.4504 | -104.6715
19.454 | -104.6727
19.4577 | -104.6724
19.4613 | -104.6777
19.465 | -104.6841
19.4687 | -104.6853
19.4723 | -104.6864
19.476 | -104.6869
19.4796 | -104.6849
19.4833 | -104.6836
19.4869 | -104.6825
19.4906 | -104.683
19.4943 | -104.6862
19.4979 | -104.686
19.5016 | -104.684
19.5052 | -104.6822
19.5089 | -104.679
19.5126 | -104.6756
19.5162 | -104.672
19.5199 | -104.6686
19.5235 | -104.6654
19.5272 | -104.6634
19.5308 | -104.6642
19.5345 | -104.6628
19.5382 | -104.6609
19.5418 | -104.6619
19.5455 | -104.6626
19.5491 | -104.6635
19.5528 | -104.6639
19.5565 | -104.6642
19.5601 | -104.6644
19.5638 | -104.6686
19.5674 | -104.6749
19.5711 | -104.682
19.5747 | -104.6929
19.5784 | -104.7047
19.5821 | -104.7133
19.5857 | -104.7189
19.5894 | -104.725
19.593 | -104.7307
19.5967 | -104.7383
19.6003 | -104.7472
19.604 | -104.7537
19.6077 | -104.7619
19.6113 | -104.7734
19.615 | -104.7838
19.6186 | -104.7932
19.6223 | -104.8038
19.626 | -104.8142
19.6296 | -104.8243
19.6333 | -104.8345
19.6369 | -104.8468
19.6406 | -104.857
19.6442 | -104.8633
19.6479 | -104.8706
19.6516 | -104.8771
19.6552 | -104.8835
19.6589 | -104.8905
19.6625 | -104.8943
19.6662 | -104.8965
19.6699 | -104.898
19.6735 | -104.9009
19.6772 | -104.9019
19.6808 | -104.8988
19.6845 | -104.8952
19.6881 | -104.8927
19.6918 | -104.8899
19.6955 | -104.8872
19.6991 | -104.884
19.7028 | -104.8775
19.7064 | -104.8691
19.7101 | -104.8587
19.7138 | -104.8495
19.7174 | -104.8426
19.7211 | -104.8351
19.7247 | -104.8242
19.7284 | -104.8132
19.732 | -104.8033
19.7357 | -104.794
19.7394 | -104.7833
19.743 | -104.7715
19.7467 | -104.7602
19.7503 | -104.7497
19.754 | -104.7407
19.7576 | -104.7312
19.7613 | -104.7213
19.765 | -104.7103
19.7686 | -104.7004
19.7723 | -104.6925
19.7759 | -104.6856
19.7796 | -104.6783
19.7833 | -104.6716
19.7869 | -104.6644
19.7906 | -104.6585
19.7942 | -104.654
19.7979 | -104.6513
19.8015 | -104.6497
19.8052 | -104.6486
19.8089 | -104.6488
19.8125 | -104.6475
19.8162 | -104.6442
19.8198 | -104.6401
19.8235 | -104.6373
19.8272 | -104.6349
19.8308 | -104.6327
19.8345 | -104.6312
19.8381 | -104.6309
19.8418 | -104.6313
19.8454 | -104.6299
19.8491 | -104.6263
19.8528 | -104.6214
19.8564 | -104.6185
19.8601 | -104.6146
19.8637 | -104.6109
19.8674 | -104.6086
19.8711 | -104.6079
19.8747 | -104.6052
19.8784 | -104.5988
19.882 | -104.5935
19.8857 | -104.5896
19.8893 | -104.5825
19.893 | -104.5734
19.8967 | -104.5641
19.9003 | -104.5535
19.904 | -104.5424
19.9076 | -104.5306
19.9113 | -104.5179
19.9149 | -104.5048
19.9186 | -104.4899
19.9223 | -104.4756
19.9259 | -104.4611
19.9296 | -104.446
19.9332 | -104.4297
19.9369 | -104.4129
19.9406 | -104.3983
19.9442 | -104.3833
19.9479 | -104.3659
19.9515 | -104.3485
19.9552 | -104.3317
19.9588 | -104.3151
19.9625 | -104.3019
19.9662 | -104.2908
19.9698 | -104.2788
19.9735 | -104.2646
19.9771 | -104.2526
19.9808 | -104.2418
19.9845 | -104.2306
19.9881 | -104.2178
19.9918 | -104.2066
19.9954 | -104.1958
19.9991 | -104.1862
20.0027 | -104.1777
20.0064 | -104.1727
20.0101 | -104.1666
20.0137 | -104.1585
20.0174 | -104.1502
20.021 | -104.1431
20.0247 | -104.139
20.0284 | -104.1318
20.032 | -104.1236
20.0357 | -104.117
20.0393 | -104.1118
20.043 | -104.1078
20.0466 | -104.1033
20.0503 | -104.0967
20.054 | -104.09
20.0576 | -104.0824
20.0613 | -104.0763
20.0649 | -104.0719
20.0686 | -104.0671
20.0722 | -104.0611
20.0759 | -104.0558
20.0796 | -104.0508
20.0832 | -104.0448
20.0869 | -104.0378
20.0905 | -104.0292
20.0942 | -104.0192
20.0979 | -104.0083
20.1015 | -104.0002
20.1052 | -103.9947
20.1088 | -103.9882
20.1125 | -103.9819
20.1161 | -103.9746
20.1198 | -103.9669
20.1235 | -103.9581
20.1271 | -103.9497
20.1308 | -103.9422
20.1344 | -103.9334
20.1381 | -103.9221
20.1418 | -103.9097
20.1454 | -103.9001
20.1491 | -103.8911
20.1527 | -103.8803
20.1564 | -103.8689
20.16 | -103.8597
20.1637 | -103.8517
20.1674 | -103.8432
20.171 | -103.8338
20.1747 | -103.8251
20.1783 | -103.8151
20.182 | -103.8058
20.1857 | -103.7989
20.1893 | -103.7893
20.193 | -103.7795
20.1966 | -103.7692
20.2003 | -103.7592
20.2039 | -103.7505
20.2076 | -103.7446
20.2113 | -103.739
20.2149 | -103.7307
20.2186 | -103.7235
20.2222 | -103.7158
20.2259 | -103.7087
20.2295 | -103.7056
20.2332 | -103.7026
20.2369 | -103.697
20.2405 | -103.6929
20.2442 | -103.6897
20.2478 | -103.6844
20.2515 | -103.6788
20.2552 | -103.6732
20.2588 | -103.6699
20.2625 | -103.6651
20.2661 | -103.6591
20.2698 | -103.6561
20.2734 | -103.6564
20.2771 | -103.6578
20.2808 | -103.6593
20.2844 | -103.6596
20.2881 | -103.6591
20.2917 | -103.6572
20.2954 | -103.6564
20.2991 | -103.6569
20.3027 | -103.6601
20.3064 | -103.6628
20.31 | -103.664
20.3137 | -103.6661
20.3173 | -103.6687
20.321 | -103.6681
20.3247 | -103.6676
20.3283 | -103.667
20.332 | -103.6674
20.3356 | -103.6681
20.3393 | -103.6688
20.343 | -103.6699
20.3466 | -103.6704
20.3503 | -103.6713
20.3539 | -103.6702
20.3576 | -103.669
20.3612 | -103.671
20.3649 | -103.6723
20.3686 | -103.6745
20.3722 | -103.6776
20.3759 | -103.6799
20.3795 | -103.6846
20.3832 | -103.6905
20.3868 | -103.6949
20.3905 | -103.699
20.3942 | -103.7035
20.3978 | -103.7047
20.4015 | -103.7059
20.4051 | -103.7121
20.4088 | -103.7162
20.4125 | -103.7207
20.4161 | -103.726
20.4198 | -103.7301
20.4234 | -103.7361
20.4271 | -103.7422
20.4307 | -103.7473
20.4344 | -103.7506
20.4381 | -103.7528
20.4417 | -103.7567
20.4454 | -103.7609
20.449 | -103.7643
20.4527 | -103.7653
20.4564 | -103.7662
20.46 | -103.7674
20.4637 | -103.7683
20.4673 | -103.7689
20.471 | -103.769
20.4746 | -103.7667
20.4783 | -103.7652
20.482 | -103.763
20.4856 | -103.7609
20.4893 | -103.7595
20.4929 | -103.7542
20.4966 | -103.7434
20.5003 | -103.7349
20.5039 | -103.7299
20.5076 | -103.723
20.5112 | -103.717
20.5149 | -103.7076
20.5185 | -103.698
20.5222 | -103.6936
20.5259 | -103.6887
20.5295 | -103.6807
20.5332 | -103.6745
20.5368 | -103.6706
20.5405 | -103.6651
20.5441 | -103.6606
20.5478 | -103.6582
20.5515 | -103.6557
20.5551 | -103.6525
20.5588 | -103.6484
20.5624 | -103.646
20.5661 | -103.6409
20.5698 | -103.633
20.5734 | -103.6291
20.5771 | -103.6273
20.5807 | -103.6236
20.5844 | -103.6223
20.588 | -103.6191
20.5917 | -103.6164
20.5954 | -103.6132
20.599 | -103.6078
20.6027 | -103.6035
20.6063 | -103.6009
20.61 | -103.5968
20.6137 | -103.5961
20.6173 | -103.5934
20.621 | -103.5905
20.6246 | -103.5884
20.6283 | -103.5864
20.6319 | -103.5857
20.6356 | -103.5845
20.6393 | -103.5817
20.6429 | -103.5798
20.6466 | -103.582
20.6502 | -103.5843
20.6539 | -103.5847
20.6576 | -103.586
20.6612 | -103.5872
20.6649 | -103.5889
20.6685 | -103.594
20.6722 | -103.5993
20.6758 | -103.603
20.6795 | -103.6093
20.6832 | -103.6157
20.6868 | -103.6219
20.6905 | -103.6296
20.6941 | -103.6382
20.6978 | -103.6445
20.7015 | -103.6535
20.7051 | -103.6646
20.7088 | -103.6755
20.7124 | -103.6879
20.7161 | -103.6991
20.7197 | -103.7121
20.7234 | -103.7251
20.7271 | -103.7371
20.7307 | -103.7508
20.7344 | -103.764
20.738 | -103.7772
20.7417 | -103.7914
20.7453 | -103.8062
20.749 | -103.8221
20.7527 | -103.8381
20.7563 | -103.8529
20.76 | -103.8702
20.7636 | -103.8859
20.7673 | -103.9004
20.771 | -103.9174
20.7746 | -103.9331
20.7783 | -103.9466
20.7819 | -103.9597
20.7856 | -103.9764
20.7892 | -103.9948
20.7929 | -104.0079
20.7966 | -104.0187
20.8002 | -104.0303
20.8039 | -104.0398
20.8075 | -104.0493
20.8112 | -104.0561
20.8149 | -104.0595
20.8185 | -104.0631
20.8222 | -104.0676
20.8258 | -104.0693
20.8295 | -104.067
20.8331 | -104.0648
20.8368 | -104.0615
20.8405 | -104.0548
20.8441 | -104.0474
20.8478 | -104.0417
20.8514 | -104.0323
20.8551 | -104.0211
20.8588 | -104.0086
20.8624 | -103.9962
20.8661 | -103.9784
20.8697 | -103.9592
20.8734 | -103.9437
20.877 | -103.9263
20.8807 | -103.9057
20.8844 | -103.887
20.888 | -103.864
20.8917 | -103.8406
20.8953 | -103.8191
20.899 | -103.7952
20.9026 | -103.7726
20.9063 | -103.7524
20.91 | -103.7293
20.9136 | -103.7073
20.9173 | -103.6843
20.9209 | -103.659
20.9246 | -103.6361
20.9283 | -103.614
20.9319 | -103.5924
20.9356 | -103.5715
20.9392 | -103.5481
20.9429 | -103.5247
20.9465 | -103.5012
20.9502 | -103.4788
20.9539 | -103.4561
20.9575 | -103.4338
20.9612 | -103.414
20.9648 | -103.3938
20.9685 | -103.3744
20.9722 | -103.357
20.9758 | -103.3413
20.9795 | -103.3251
20.9831 | -103.3106
20.9868 | -103.2988
20.9904 | -103.2884
20.9941 | -103.2769
20.9978 | -103.2635
21.0014 | -103.2519
21.0051 | -103.2419
21.0087 | -103.232
21.0124 | -103.2258
21.016 | -103.2221
21.0197 | -103.2175
21.0234 | -103.2161
21.027 | -103.2158
21.0307 | -103.2151
21.0343 | -103.2142
21.038 | -103.2143
21.0417 | -103.2158
21.0453 | -103.2175
21.049 | -103.2185
21.0526 | -103.22
21.0563 | -103.2216
21.0599 | -103.2252
21.0636 | -103.23
21.0673 | -103.2334
21.0709 | -103.2393
21.0746 | -103.2462
21.0782 | -103.2532
21.0819 | -103.256
21.0856 | -103.2576
21.0892 | -103.2616
21.0929 | -103.2648
21.0965 | -103.2641
21.1002 | -103.2621
21.1038 | -103.263
21.1075 | -103.2607
21.1112 | -103.2629
21.1148 | -103.2631
21.1185 | -103.2647
21.1221 | -103.2656
21.1258 | -103.2652
21.1295 | -103.2635
21.1331 | -103.2593
21.1368 | -103.2565
21.1404 | -103.2543
21.1441 | -103.2521
21.1477 | -103.2472
21.1514 | -103.2423
21.1551 | -103.2353
21.1587 | -103.2305
21.1624 | -103.2262
21.166 | -103.2215
21.1697 | -103.2126
21.1733 | -103.2004
21.177 | -103.19
21.1807 | -103.181
21.1843 | -103.1701
21.188 | -103.1581
21.1916 | -103.1477
21.1953 | -103.136
21.199 | -103.1228
21.2026 | -103.1087
21.2063 | -103.0935
21.2099 | -103.0776
21.2136 | -103.063
21.2172 | -103.0491
21.2209 | -103.0376
21.2246 | -103.0246
21.2282 | -103.0108
21.2319 | -102.9964
21.2355 | -102.9827
21.2392 | -102.9722
21.2429 | -102.9611
21.2465 | -102.9521
21.2502 | -102.9441
21.2538 | -102.9363
21.2575 | -102.9289
21.2611 | -102.9218
21.2648 | -102.9157
21.2685 | -102.9099
21.2721 | -102.9049
21.2758 | -102.898
21.2794 | -102.8927
21.2831 | -102.888
21.2868 | -102.8821
21.2904 | -102.8774
21.2941 | -102.8758
21.2977 | -102.8759
21.3014 | -102.8762
21.305 | -102.8745
21.3087 | -102.8742
21.3124 | -102.8752
21.316 | -102.8741
21.3197 | -102.8744
21.3233 | -102.8736
21.327 | -102.8733
21.3307 | -102.8761
21.3343 | -102.8796
21.338 | -102.8824
21.3416 | -102.8855
21.3453 | -102.8892
21.3489 | -102.8932
21.3526 | -102.8989
21.3563 | -102.9033
21.3599 | -102.9063
21.3636 | -102.9111
21.3672 | -102.9157
21.3709 | -102.9176
21.3745 | -102.9197
21.3782 | -102.9201
21.3819 | -102.9227
21.3855 | -102.9262
21.3892 | -102.9295
21.3928 | -102.9349
21.3965 | -102.943
21.4002 | -102.9507
21.4038 | -102.9554
21.4075 | -102.96
21.4111 | -102.9604
21.4148 | -102.9615
21.4184 | -102.9659
21.4221 | -102.9673
21.4258 | -102.9673
21.4294 | -102.9692
21.4331 | -102.9731
21.4367 | -102.9787
21.4404 | -102.9817
21.4441 | -102.9857
21.4477 | -102.9906
21.4514 | -102.9963
21.455 | -102.9996
21.4587 | -103.0005
21.4623 | -103.0022
21.466 | -103.0044
21.4697 | -103.0058
21.4733 | -103.011
21.477 | -103.0163
21.4806 | -103.0201
21.4843 | -103.0236
21.488 | -103.028
21.4916 | -103.0337
21.4953 | -103.0367
21.4989 | -103.0416
21.5026 | -103.0482
21.5062 | -103.0532
21.5099 | -103.0585
21.5136 | -103.0634
21.5172 | -103.0674
21.5209 | -103.0707
21.5245 | -103.0722
21.5282 | -103.0753
21.5318 | -103.0792
21.5355 | -103.0837
21.5392 | -103.0904
21.5428 | -103.0936
21.5465 | -103.095
21.5501 | -103.0965
21.5538 | -103.097
21.5575 | -103.0971
21.5611 | -103.0994
21.5648 | -103.1004
21.5684 | -103.0994
21.5721 | -103.101
21.5757 | -103.1041
21.5794 | -103.1076
21.5831 | -103.1118
21.5867 | -103.1157
21.5904 | -103.1199
21.594 | -103.1211
21.5977 | -103.1193
21.6014 | -103.1199
21.605 | -103.1225
21.6087 | -103.1271
21.6123 | -103.1306
21.616 | -103.1316
21.6196 | -103.1348
21.6233 | -103.1384
21.627 | -103.1392
21.6306 | -103.1387
21.6343 | -103.1378
21.6379 | -103.1388
21.6416 | -103.1392
21.6453 | -103.1416
21.6489 | -103.1428
21.6526 | -103.1409
21.6562 | -103.1399
21.6599 | -103.1389
21.6635 | -103.1377
21.6672 | -103.1359
21.6709 | -103.1328
21.6745 | -103.131
21.6782 | -103.1304
21.6818 | -103.1271
21.6855 | -103.126
21.6891 | -103.1246
21.6928 | -103.1227
21.6965 | -103.1212
21.7001 | -103.1175
21.7038 | -103.1126
21.7074 | -103.1074
21.7111 | -103.1007
21.7148 | -103.093
21.7184 | -103.091
21.7221 | -103.0871
21.7257 | -103.0798
21.7294 | -103.075
21.733 | -103.0719
21.7367 | -103.0694
21.7404 | -103.066
21.744 | -103.0609
21.7477 | -103.0571
21.7513 | -103.0521
21.755 | -103.0444
21.7587 | -103.0357
21.7623 | -103.0322
21.766 | -103.0326
21.7696 | -103.0287
21.7733 | -103.0242
21.7769 | -103.0207
21.7806 | -103.0152
21.7843 | -103.0088
21.7879 | -102.9995
21.7916 | -102.9894
21.7952 | -102.9813
21.7989 | -102.9735
21.8026 | -102.9637
21.8062 | -102.9541
21.8099 | -102.9464
21.8135 | -102.9379
21.8172 | -102.9264
21.8208 | -102.9177
21.8245 | -102.9102
21.8282 | -102.9025
21.8318 | -102.8951
21.8355 | -102.8862
21.8391 | -102.8757
21.8428 | -102.8665
21.8464 | -102.8578
21.8501 | -102.849
21.8538 | -102.8408
21.8574 | -102.8337
21.8611 | -102.8282
21.8647 | -102.8235
21.8684 | -102.8158
21.8721 | -102.8095
21.8757 | -102.8049
21.8794 | -102.7977
21.883 | -102.7924
21.8867 | -102.789
21.8903 | -102.7845
21.894 | -102.7807
21.8977 | -102.7794
21.9013 | -102.7802
21.905 | -102.7843
21.9086 | -102.7863
21.9123 | -102.7853
21.916 | -102.7833
21.9196 | -102.781
21.9233 | -102.7795
21.9269 | -102.7797
21.9306 | -102.7816
21.9342 | -102.7859
21.9379 | -102.7887
21.9416 | -102.7898
21.9452 | -102.7893
21.9489 | -102.7916
21.9525 | -102.7947
21.9562 | -102.7969
21.9599 | -102.8016
21.9635 | -102.8062
21.9672 | -102.8101
21.9708 | -102.8113
21.9745 | -102.8146
21.9781 | -102.8197
21.9818 | -102.8249
21.9855 | -102.8278
21.9891 | -102.8329
21.9928 | -102.8382
21.9964 | -102.8406
22.0001 | -102.8444
22.0038 | -102.8488
22.0074 | -102.8514
22.0111 | -102.8548
22.0147 | -102.8579
22.0184 | -102.8625
22.022 | -102.8687
22.0257 | -102.875
22.0294 | -102.8782
22.033 | -102.882
22.0367 | -102.8873
22.0403 | -102.8885
22.044 | -102.8898
22.0476 | -102.8913
22.0513 | -102.8938
22.055 | -102.8975
22.0586 | -102.901
22.0623 | -102.9047
22.0659 | -102.9105
22.0696 | -102.9161
22.0733 | -102.9187
22.0769 | -102.9192
22.0806 | -102.9209
22.0842 | -102.9214
22.0879 | -102.9194
22.0915 | -102.9198
22.0952 | -102.9203
22.0989 | -102.9229
22.1025 | -102.9256
22.1062 | -102.9256
22.1098 | -102.9261
22.1135 | -102.9244
22.1172 | -102.9205
22.1208 | -102.9136
22.1245 | -102.9086
22.1281 | -102.906
22.1318 | -102.901
22.1354 | -102.8939
22.1391 | -102.8863
22.1428 | -102.8773
22.1464 | -102.8661
22.1501 | -102.8571
22.1537 | -102.8489
22.1574 | -102.8389
22.161 | -102.8305
22.1647 | -102.8157
22.1684 | -102.8064
22.172 | -102.7954
22.1757 | -102.7824
22.1793 | -102.7664
22.183 | -102.7512
22.1867 | -102.7383
22.1903 | -102.7249
22.194 | -102.7125
22.1976 | -102.6985
22.2013 | -102.6823
22.2049 | -102.6671
22.2086 | -102.6517
22.2123 | -102.6378
22.2159 | -102.6247
22.2196 | -102.6102
22.2232 | -102.5978
22.2269 | -102.5872
22.2306 | -102.5788
22.2342 | -102.5719
22.2379 | -102.5647
22.2415 | -102.5564
22.2452 | -102.5475
22.2488 | -102.5383
22.2525 | -102.5316
22.2562 | -102.525
22.2598 | -102.5214
22.2635 | -102.5178
22.2671 | -102.5152
22.2708 | -102.5129
22.2745 | -102.5088
22.2781 | -102.509
22.2818 | -102.5111
22.2854 | -102.514
22.2891 | -102.5134
22.2927 | -102.5147
22.2964 | -102.5182
22.3001 | -102.5243
22.3037 | -102.5303
22.3074 | -102.5384
22.311 | -102.5485
22.3147 | -102.5591
22.3183 | -102.5706
22.322 | -102.5815
22.3257 | -102.5924
22.3293 | -102.6055
22.333 | -102.6203
22.3366 | -102.6346
22.3403 | -102.6499
22.344 | -102.6668
22.3476 | -102.685
22.3513 | -102.7032
22.3549 | -102.7247
22.3586 | -102.7451
22.3622 | -102.7636
22.3659 | -102.7812
22.3696 | -102.7991
22.3732 | -102.8175
22.3769 | -102.8364
22.3805 | -102.8538
22.3842 | -102.8712
22.3879 | -102.8844
22.3915 | -102.8982
22.3952 | -102.9162
22.3988 | -102.9341
22.4025 | -102.9504
22.4061 | -102.9626
22.4098 | -102.9779
22.4135 | -102.9908
22.4171 | -102.999
22.4208 | -103.0057
22.4244 | -103.0136
22.4281 | -103.0243
22.4318 | -103.0332
22.4354 | -103.0406
22.4391 | -103.0443
22.4427 | -103.0471
22.4464 | -103.0512
22.45 | -103.0514
22.4537 | -103.0439
22.4574 | -103.031
22.461 | -103.0061
22.4647 | -102.9611
22.4683 | -102.9206
22.472 | -102.8887
22.4757 | -102.8611
22.4793 | -102.8327
22.483 | -102.8027
22.4866 | -102.7696
22.4903 | -102.7324
22.4939 | -102.6921
22.4976 | -102.6518
22.5013 | -102.6086
22.5049 | -102.5615
22.5086 | -102.5123
22.5122 | -102.4612
22.5159 | -102.4092
22.5195 | -102.3572
22.5232 | -102.3005
22.5269 | -102.2416
22.5305 | -102.1809
22.5342 | -102.1206
22.5378 | -102.0599
22.5415 | -101.9988
22.5452 | -101.9372
22.5488 | -101.8783
22.5525 | -101.8219
22.5561 | -101.7661
22.5598 | -101.7115
22.5634 | -101.6587
22.5671 | -101.6089
22.5708 | -101.5608
22.5744 | -101.517
22.5781 | -101.4744
22.5817 | -101.4338
22.5854 | -101.3962
22.5891 | -101.3628
22.5927 | -101.3335
22.5964 | -101.3069
22.6 | -101.2843
22.6037 | -101.2671
22.6073 | -101.256
22.611 | -101.2476
22.6147 | -101.2426
22.6183 | -101.2406
22.622 | -101.2428
22.6256 | -101.251
22.6293 | -101.2631
22.633 | -101.2785
22.6366 | -101.2986
22.6403 | -101.3219
22.6439 | -101.3501
22.6476 | -101.3829
22.6512 | -101.4178
22.6549 | -101.4571
22.6586 | -101.5007
22.6622 | -101.5462
22.6659 | -101.5941
22.6695 | -101.6453
22.6732 | -101.6996
22.6768 | -101.7566
22.6805 | -101.816
22.6842 | -101.8769
22.6878 | -101.9401
22.6915 | -102.0043
22.6951 | -102.0659
22.6988 | -102.1281
22.7025 | -102.1924
22.7061 | -102.2565
22.7098 | -102.3229
22.7134 | -102.3889
22.7171 | -102.4502
22.7207 | -102.51
22.7244 | -102.5715
22.7281 | -102.6327
22.7317 | -102.6885
22.7354 | -102.744
22.739 | -102.8017
22.7427 | -102.8577
22.7464 | -102.9085
22.75 | -102.9556
22.7537 | -103.0021
22.7573 | -103.0479
22.761 | -103.0914
22.7646 | -103.1307
22.7683 | -103.1658
22.772 | -103.2021
22.7756 | -103.2415
22.7793 | -103.2822
22.7829 | -103.3307
22.7866 | -103.3881
22.7903 | -103.4275
22.7939 | -103.4457
22.7976 | -103.455
22.8012 | -103.4656
22.8049 | -103.4751
22.8085 | -103.4816
22.8122 | -103.488
22.8159 | -103.4969
22.8195 | -103.5048
22.8232 | -103.5109
22.8268 | -103.5119
22.8305 | -103.5114
22.8341 | -103.51
22.8378 | -103.5051
22.8415 | -103.4995
22.8451 | -103.4962
22.8488 | -103.4949
22.8524 | -103.4911
22.8561 | -103.484
22.8598 | -103.4784
22.8634 | -103.472
22.8671 | -103.4634
22.8707 | -103.4503
22.8744 | -103.4364
22.878 | -103.4232
22.8817 | -103.411
22.8854 | -103.3975
22.889 | -103.385
22.8927 | -103.3727
22.8963 | -103.3582
22.9 | -103.3433
22.9037 | -103.3277
22.9073 | -103.3114
22.911 | -103.2975
22.9146 | -103.2851
22.9183 | -103.269
22.9219 | -103.2552
22.9256 | -103.241
22.9293 | -103.227
22.9329 | -103.212
22.9366 | -103.1976
22.9402 | -103.1845
22.9439 | -103.172
22.9476 | -103.159
22.9512 | -103.1469
22.9549 | -103.1385
22.9585 | -103.1308
22.9622 | -103.1247
22.9658 | -103.119
22.9695 | -103.1124
22.9732 | -103.1067
22.9768 | -103.1005
22.9805 | -103.0938
22.9841 | -103.087
22.9878 | -103.0808
22.9914 | -103.0779
22.9951 | -103.0737
22.9988 | -103.0718
23.0024 | -103.0711
23.0061 | -103.0705
23.0097 | -103.0698
23.0134 | -103.0681
23.0171 | -103.0681
23.0207 | -103.067
23.0244 | -103.065
23.028 | -103.064
23.0317 | -103.0655
23.0353 | -103.0693
23.039 | -103.0713
23.0427 | -103.0728
23.0463 | -103.0759
23.05 | -103.0769
23.0536 | -103.077
23.0573 | -103.0791
23.061 | -103.0798
23.0646 | -103.0791
23.0683 | -103.0784
23.0719 | -103.0785
23.0756 | -103.0795
23.0792 | -103.0792
23.0829 | -103.0795
23.0866 | -103.0791
23.0902 | -103.0771
23.0939 | -103.0784
23.0975 | -103.0772
23.1012 | -103.0763
23.1049 | -103.0754
23.1085 | -103.0739
23.1122 | -103.0705
23.1158 | -103.0641
23.1195 | -103.0619
23.1231 | -103.0596
23.1268 | -103.0539
23.1305 | -103.0483
23.1341 | -103.0442
23.1378 | -103.0386
23.1414 | -103.0307
23.1451 | -103.0239
23.1487 | -103.0168
23.1524 | -103.0102
23.1561 | -103.0053
23.1597 | -102.9997
23.1634 | -102.9929
23.167 | -102.9841
23.1707 | -102.9747
23.1744 | -102.9651
23.178 | -102.9543
23.1817 | -102.9417
23.1853 | -102.9288
23.189 | -102.9178
23.1926 | -102.9083
23.1963 | -102.9006
23.2 | -102.891
23.2036 | -102.8815
23.2073 | -102.8715
23.2109 | -102.8631
23.2146 | -102.8542
23.2183 | -102.8484
23.2219 | -102.8397
23.2256 | -102.8331
23.2292 | -102.8276
23.2329 | -102.8234
23.2365 | -102.8186
23.2402 | -102.8132
23.2439 | -102.8098
23.2475 | -102.8091
23.2512 | -102.8069
23.2548 | -102.8056
23.2585 | -102.8033
23.2622 | -102.8029
23.2658 | -102.8033
23.2695 | -102.8039
23.2731 | -102.8032
23.2768 | -102.8006
23.2804 | -102.8009
23.2841 | -102.8023
23.2878 | -102.805
23.2914 | -102.8087
23.2951 | -102.8102
23.2987 | -102.813
23.3024 | -102.818
23.306 | -102.8226
23.3097 | -102.8259
23.3134 | -102.8304
23.317 | -102.8345
23.3207 | -102.8377
23.3243 | -102.8402
23.328 | -102.8406
23.3317 | -102.8427
23.3353 | -102.8466
23.339 | -102.8485
23.3426 | -102.8533
23.3463 | -102.8551
23.3499 | -102.8586
23.3536 | -102.8623
23.3573 | -102.862
23.3609 | -102.8597
23.3646 | -102.8575
23.3682 | -102.8549
23.3719 | -102.8524
23.3756 | -102.8493
23.3792 | -102.8472
23.3829 | -102.8444
23.3865 | -102.841
23.3902 | -102.8392
23.3938 | -102.8375
23.3975 | -102.8353
23.4012 | -102.8318
23.4048 | -102.8265
23.4085 | -102.8216
23.4121 | -102.8178
23.4158 | -102.8144
23.4195 | -102.8113
23.4231 | -102.806
23.4268 | -102.8012
23.4304 | -102.7956
23.4341 | -102.7898
23.4377 | -102.7841
23.4414 | -102.782
23.4451 | -102.7817
23.4487 | -102.7802
23.4524 | -102.7781
23.456 | -102.7766
23.4597 | -102.7755
23.4633 | -102.7747
23.467 | -102.7775
23.4707 | -102.7804
23.4743 | -102.7825
23.478 | -102.7847
23.4816 | -102.7866
23.4853 | -102.7863
23.489 | -102.7891
23.4926 | -102.7943
23.4963 | -102.8016
23.4999 | -102.8094
23.5036 | -102.8154
23.5072 | -102.8231
23.5109 | -102.834
23.5146 | -102.8444
23.5182 | -102.8529
23.5219 | -102.8624
23.5255 | -102.8717
23.5292 | -102.8821
23.5329 | -102.8929
23.5365 | -102.9054
23.5402 | -102.9175
23.5438 | -102.9297
23.5475 | -102.9451
23.5511 | -102.9615
23.5548 | -102.9756
23.5585 | -102.9896
23.5621 | -103.0049
23.5658 | -103.0206
23.5694 | -103.0384
23.5731 | -103.0555
23.5768 | -103.0724
23.5804 | -103.0896
23.5841 | -103.1068
23.5877 | -103.1252
23.5914 | -103.1458
23.595 | -103.1654
23.5987 | -103.1858
23.6024 | -103.2084
23.606 | -103.232
23.6097 | -103.2578
23.6133 | -103.2784
23.617 | -103.2966
23.6207 | -103.3147
23.6243 | -103.3344
23.628 | -103.3549
23.6316 | -103.3763
23.6353 | -103.3949
23.6389 | -103.4123
23.6426 | -103.4302
23.6463 | -103.4465
23.6499 | -103.4634
23.6536 | -103.4771
23.6572 | -103.4885
23.6609 | -103.4992
23.6645 | -103.5089
23.6682 | -103.5187
23.6719 | -103.5283
23.6755 | -103.5345
23.6792 | -103.5421
23.6828 | -103.5484
23.6865 | -103.5538
23.6902 | -103.5581
23.6938 | -103.5617
23.6975 | -103.5627
23.7011 | -103.5626
23.7048 | -103.5605
23.7084 | -103.5586
23.7121 | -103.5553
23.7158 | -103.5509
23.7194 | -103.5441
23.7231 | -103.5353
23.7267 | -103.5258
23.7304 | -103.5164
23.7341 | -103.5068
23.7377 | -103.4993
23.7414 | -103.4902
23.745 | -103.4769
23.7487 | -103.4627
23.7523 | -103.4503
23.756 | -103.4393
23.7597 | -103.4262
23.7633 | -103.4152
23.767 | -103.4047
23.7706 | -103.3935
23.7743 | -103.3833
23.778 | -103.372
23.7816 | -103.3616
23.7853 | -103.3514
23.7889 | -103.3405
23.7926 | -103.3274
23.7962 | -103.3155
23.7999 | -103.3055
23.8036 | -103.294
23.8072 | -103.2842
23.8109 | -103.2738
23.8145 | -103.2632
23.8182 | -103.2543
23.8218 | -103.2473
23.8255 | -103.2403
23.8292 | -103.233
23.8328 | -103.2261
23.8365 | -103.2187
23.8401 | -103.2126
23.8438 | -103.2066
23.8475 | -103.2019
23.8511 | -103.1975
23.8548 | -103.1942
23.8584 | -103.1921
23.8621 | -103.1927
23.8657 | -103.1934
23.8694 | -103.1939
23.8731 | -103.1946
23.8767 | -103.1952
23.8804 | -103.1956
23.884 | -103.196
23.8877 | -103.1977
23.8914 | -103.2027
23.895 | -103.2051
23.8987 | -103.2065
23.9023 | -103.2091
23.906 | -103.2122
23.9096 | -103.2162
23.9133 | -103.22
23.917 | -103.225
23.9206 | -103.2295
23.9243 | -103.2332
23.9279 | -103.2365
23.9316 | -103.2422
23.9353 | -103.2478
23.9389 | -103.2522
23.9426 | -103.2577
23.9462 | -103.2634
23.9499 | -103.2702
23.9535 | -103.2755
23.9572 | -103.2803
23.9609 | -103.2838
23.9645 | -103.2883
23.9682 | -103.2913
23.9718 | -103.2937
23.9755 | -103.2974
23.9791 | -103.3017
23.9828 | -103.3048
23.9865 | -103.306
23.9901 | -103.3077
23.9938 | -103.3127
23.9974 | -103.3169
24.0011 | -103.3166
24.0048 | -103.3144
24.0084 | -103.3145
24.0121 | -103.3128
24.0157 | -103.3095
24.0194 | -103.3075
24.023 | -103.3058
24.0267 | -103.3033
24.0304 | -103.2991
24.034 | -103.294
24.0377 | -103.2885
24.0413 | -103.2818
24.045 | -103.2732
24.0487 | -103.2659
24.0523 | -103.2581
24.056 | -103.2494
24.0596 | -103.2398
24.0633 | -103.2306
24.0669 | -103.2208
24.0706 | -103.2108
24.0743 | -103.1989
24.0779 | -103.1853
24.0816 | -103.1723
24.0852 | -103.1594
24.0889 | -103.1479
24.0926 | -103.133
24.0962 | -103.1145
24.0999 | -103.0988
24.1035 | -103.0819
24.1072 | -103.0644
24.1108 | -103.0478
24.1145 | -103.0321
24.1182 | -103.0146
24.1218 | -102.9965
24.1255 | -102.9802
24.1291 | -102.9644
24.1328 | -102.9492
24.1364 | -102.9352
24.1401 | -102.9213
24.1438 | -102.9084
24.1474 | -102.8948
24.1511 | -102.8828
24.1547 | -102.8733
24.1584 | -102.863
24.1621 | -102.8512
24.1657 | -102.8393
24.1694 | -102.83
24.173 | -102.8208
24.1767 | -102.8126
24.1803 | -102.8053
24.184 | -102.7995
24.1877 | -102.7943
24.1913 | -102.7901
24.195 | -102.786
24.1986 | -102.782
24.2023 | -102.7781
24.206 | -102.775
24.2096 | -102.7746
24.2133 | -102.7753
24.2169 | -102.7772
24.2206 | -102.7799
24.2242 | -102.7827
24.2279 | -102.786
24.2316 | -102.7894
24.2352 | -102.7951
24.2389 | -102.8023
24.2425 | -102.81
24.2462 | -102.8181
24.2499 | -102.826
24.2535 | -102.8349
24.2572 | -102.8432
24.2608 | -102.8513
24.2645 | -102.8558
24.2681 | -102.8442
24.2718 | -102.7933
24.2755 | -102.5332
24.2791 | -102.3056
24.2828 | -102.0839
24.2864 | -101.9295
24.2901 | -101.8001
24.2937 | -101.6854
24.2974 | -101.5687
24.3011 | -101.4472
24.3047 | -101.3213
24.3084 | -101.1888
24.312 | -101.0503
24.3157 | -100.9074
24.3194 | -100.7582
24.323 | -100.6025
24.3267 | -100.4445
24.3303 | -100.2819
24.334 | -100.1163
24.3376 | -99.9486
24.3413 | -99.7776
24.345 | -99.6061
24.3486 | -99.4353
24.3523 | -99.2655
24.3559 | -99.0967
24.3596 | -98.9289
24.3633 | -98.7627
24.3669 | -98.5995
24.3706 | -98.4408
24.3742 | -98.2862
24.3779 | -98.1356
24.3815 | -97.9901
24.3852 | -97.8497
24.3889 | -97.7145
24.3925 | -97.5846
24.3962 | -97.4605
24.3998 | -97.3431
24.4035 | -97.2327
24.4072 | -97.1286
24.4108 | -97.0317
24.4145 | -96.9424
24.4181 | -96.8607
24.4218 | -96.7873
24.4254 | -96.7218
24.4291 | -96.6638
24.4328 | -96.6135
24.4364 | -96.5718
24.4401 | -96.5392
24.4437 | -96.5143
24.4474 | -96.4974
24.451 | -96.4887
24.4547 | -96.4879
24.4584 | -96.495
24.462 | -96.5107
24.4657 | -96.5349
24.4693 | -96.5666
24.473 | -96.6059
24.4767 | -96.6531
24.4803 | -96.7079
24.484 | -96.7702
24.4876 | -96.8402
24.4913 | -96.9179
24.4949 | -97.0023
24.4986 | -97.094
24.5023 | -97.1923
24.5059 | -97.2974
24.5096 | -97.4086
24.5132 | -97.525
24.5169 | -97.6473
24.5206 | -97.7751
24.5242 | -97.9077
24.5279 | -98.045
24.5315 | -98.1855
24.5352 | -98.3299
24.5388 | -98.478
24.5425 | -98.629
24.5462 | -98.7814
24.5498 | -98.9358
24.5535 | -99.0914
24.5571 | -99.2476
24.5608 | -99.4036
24.5645 | -99.5573
24.5681 | -99.7074
24.5718 | -99.8501
24.5754 | -99.9832
24.5791 | -100.1019
24.5827 | -100.1973
24.5864 | -100.2875
24.5901 | -100.3739
24.5937 | -100.4625
24.5974 | -100.5507
24.601 | -100.6326
24.6047 | -100.7055
24.6084 | -100.7699
24.612 | -100.8309
24.6157 | -100.9013
24.6193 | -100.9876
24.623 | -101.0831
24.6266 | -101.2019
24.6303 | -101.2856
24.634 | -101.2861
24.6376 | -101.2436
24.6413 | -101.1752
24.6449 | -101.1005
24.6486 | -101.0221
24.6522 | -100.9407
24.6559 | -100.8576
24.6596 | -100.7714
24.6632 | -100.6837
24.6669 | -100.596
24.6705 | -100.5098
24.6742 | -100.422
24.6779 | -100.3338
24.6815 | -100.2476
24.6852 | -100.1616
24.6888 | -100.0789
24.6925 | -99.9985
24.6961 | -99.9195
24.6998 | -99.8419
24.7035 | -99.7677
24.7071 | -99.6976
24.7108 | -99.6306
24.7144 | -99.5679
24.7181 | -99.5084
24.7218 | -99.4532
24.7254 | -99.4027
24.7291 | -99.3566
24.7327 | -99.3164
24.7364 | -99.2812
24.74 | -99.2514
24.7437 | -99.2266
24.7474 | -99.2079
24.751 | -99.1943
24.7547 | -99.186
24.7583 | -99.184
24.762 | -99.1867
24.7657 | -99.195
24.7693 | -99.2082
24.773 | -99.2265
24.7766 | -99.251
24.7803 | -99.2815
24.7839 | -99.3174
24.7876 | -99.3577
24.7913 | -99.4036
24.7949 | -99.4553
24.7986 | -99.5108
24.8022 | -99.6002
24.8059 | -99.6622
24.8095 | -99.728
24.8132 | -99.7974
24.8169 | -99.8702
24.8205 | -99.947
24.8242 | -100.0283
24.8278 | -100.1114
24.8315 | -100.196
24.8352 | -100.2837
24.8388 | -100.3733
24.8425 | -100.4637
24.8461 | -100.5565
24.8498 | -100.6502
24.8534 | -100.7442
24.8571 | -100.8376
24.8608 | -100.9309
24.8644 | -101.0249
24.8681 | -101.1189
24.8717 | -101.2114
24.8754 | -101.3027
24.8791 | -101.3925
24.8827 | -101.4805
24.8864 | -101.5662
24.89 | -101.6488
24.8937 | -101.7298
24.8973 | -101.8085
24.901 | -101.8844
24.9047 | -101.9569
24.9083 | -102.0266
24.912 | -102.0943
24.9156 | -102.1611
24.9193 | -102.2261
24.923 | -102.2942
24.9266 | -102.3666
24.9303 | -102.4461
24.9339 | -102.5416
24.9376 | -102.6194
24.9412 | -102.6803
24.9449 | -102.7151
24.9486 | -102.7293
24.9522 | -102.7391
24.9559 | -102.7495
24.9595 | -102.7608
24.9632 | -102.7745
24.9668 | -102.786
24.9705 | -102.7927
24.9742 | -102.7989
24.9778 | -102.8076
24.9815 | -102.8159
24.9851 | -102.8243
24.9888 | -102.8358
24.9925 | -102.847
24.9961 | -102.8529
24.9998 | -102.8548
25.0034 | -102.8562
25.0071 | -102.8582
25.0107 | -102.8607
25.0144 | -102.8623
25.0181 | -102.8638
25.0217 | -102.8646
25.0254 | -102.8673
25.029 | -102.8707
25.0327 | -102.8698
25.0364 | -102.8692
25.04 | -102.8696
25.0437 | -102.8696
25.0473 | -102.8702
25.051 | -102.8699
25.0546 | -102.8701
25.0583 | -102.8707
25.062 | -102.8724
25.0656 | -102.872
25.0693 | -102.8726
25.0729 | -102.8739
25.0766 | -102.8742
25.0803 | -102.8754
25.0839 | -102.875
25.0876 | -102.8751
25.0912 | -102.8748
25.0949 | -102.8743
25.0985 | -102.8716
25.1022 | -102.8711
25.1059 | -102.8709
25.1095 | -102.8681
25.1132 | -102.8674
25.1168 | -102.8663
25.1205 | -102.8638
25.1241 | -102.8638
25.1278 | -102.8624
25.1315 | -102.8609
25.1351 | -102.8594
25.1388 | -102.8566
25.1424 | -102.8536
25.1461 | -102.8506
25.1498 | -102.8482
25.1534 | -102.8486
25.1571 | -102.8465
25.1607 | -102.8431
25.1644 | -102.8404
25.168 | -102.8389
25.1717 | -102.8353
25.1754 | -102.8282
25.179 | -102.8199
25.1827 | -102.8148
25.1863 | -102.8108
25.19 | -102.806
25.1937 | -102.8012
25.1973 | -102.7954
25.201 | -102.7881
25.2046 | -102.7778
25.2083 | -102.7705
25.2119 | -102.7629
25.2156 | -102.7548
25.2193 | -102.7449
25.2229 | -102.735
25.2266 | -102.7257
25.2302 | -102.715
25.2339 | -102.7041
25.2376 | -102.6949
25.2412 | -102.6844
25.2449 | -102.6742
25.2485 | -102.6644
25.2522 | -102.6538
25.2558 | -102.6443
25.2595 | -102.6346
25.2632 | -102.6245
25.2668 | -102.6136
25.2705 | -102.6023
25.2741 | -102.5919
25.2778 | -102.5826
25.2814 | -102.5737
25.2851 | -102.5665
25.2888 | -102.5587
25.2924 | -102.5506
25.2961 | -102.5428
25.2997 | -102.5363
25.3034 | -102.5326
25.3071 | -102.5284
25.3107 | -102.5213
25.3144 | -102.5152
25.318 | -102.5102
25.3217 | -102.5062
25.3253 | -102.5061
25.329 | -102.514
25.3327 | -102.5149
25.3363 | -102.5152
25.34 | -102.5158
25.3436 | -102.5177
25.3473 | -102.521
25.351 | -102.5259
25.3546 | -102.5332
25.3583 | -102.5417
25.3619 | -102.5507
25.3656 | -102.5578
25.3692 | -102.5674
25.3729 | -102.5783
25.3766 | -102.5886
25.3802 | -102.5998
25.3839 | -102.6106
25.3875 | -102.6244
25.3912 | -102.64
25.3949 | -102.6555
25.3985 | -102.6682
25.4022 | -102.6813
25.4058 | -102.6951
25.4095 | -102.7088
25.4131 | -102.7226
25.4168 | -102.7368
25.4205 | -102.749
25.4241 | -102.7623
25.4278 | -102.7769
25.4314 | -102.7923
25.4351 | -102.8073
25.4387 | -102.819
25.4424 | -102.8314
25.4461 | -102.8447
25.4497 | -102.8582
25.4534 | -102.8734
25.457 | -102.8877
25.4607 | -102.8976
25.4644 | -102.9067
25.468 | -102.9168
25.4717 | -102.9229
25.4753 | -102.9302
25.479 | -102.9373
25.4826 | -102.9395
25.4863 | -102.9405
25.49 | -102.9407
25.4936 | -102.9391
25.4973 | -102.9366
25.5009 | -102.9326
25.5046 | -102.9302
25.5083 | -102.9295
25.5119 | -102.925
25.5156 | -102.9185
25.5192 | -102.9117
25.5229 | -102.9014
25.5265 | -102.8903
25.5302 | -102.8826
25.5339 | -102.8764
25.5375 | -102.8675
25.5412 | -102.8549
25.5448 | -102.8412
25.5485 | -102.8276
25.5522 | -102.8133
25.5558 | -102.8002
25.5595 | -102.7859
25.5631 | -102.772
25.5668 | -102.7574
25.5704 | -102.7406
25.5741 | -102.7232
25.5778 | -102.7081
25.5814 | -102.6913
25.5851 | -102.6741
25.5887 | -102.6584
25.5924 | -102.6433
25.596 | -102.6286
25.5997 | -102.6149
25.6034 | -102.6018
25.607 | -102.5882
25.6107 | -102.5753
25.6143 | -102.5612
25.618 | -102.5477
25.6217 | -102.5341
25.6253 | -102.5217
25.629 | -102.511
25.6326 | -102.5002
25.6363 | -102.4913
25.6399 | -102.4831
25.6436 | -102.4739
25.6473 | -102.4638
25.6509 | -102.454
25.6546 | -102.4448
25.6582 | -102.4379
25.6619 | -102.4309
25.6656 | -102.4233
25.6692 | -102.4187
25.6729 | -102.4133
25.6765 | -102.4078
25.6802 | -102.4021
25.6838 | -102.3984
25.6875 | -102.3959
25.6912 | -102.3928
25.6948 | -102.389
25.6985 | -102.388
25.7021 | -102.3851
25.7058 | -102.3796
25.7095 | -102.3745
25.7131 | -102.3713
25.7168 | -102.37
25.7204 | -102.3702
25.7241 | -102.3707
25.7277 | -102.3709
25.7314 | -102.3683
25.7351 | -102.3687
25.7387 | -102.3715
25.7424 | -102.3718
25.746 | -102.3692
25.7497 | -102.3656
25.7533 | -102.3628
25.757 | -102.3595
25.7607 | -102.3577
25.7643 | -102.3566
25.768 | -102.3536
25.7716 | -102.3523
25.7753 | -102.3512
25.779 | -102.3472
25.7826 | -102.3448
25.7863 | -102.3413
25.7899 | -102.3356
25.7936 | -102.3287
25.7972 | -102.3221
25.8009 | -102.3165
25.8046 | -102.3104
25.8082 | -102.3019
25.8119 | -102.2933
25.8155 | -102.2844
25.8192 | -102.2738
25.8229 | -102.2633
25.8265 | -102.2528
25.8302 | -102.2407
25.8338 | -102.2307
25.8375 | -102.2208
25.8411 | -102.2108
25.8448 | -102.2014
25.8485 | -102.1909
25.8521 | -102.1798
25.8558 | -102.1659
25.8594 | -102.1464
25.8631 | -102.1343
25.8668 | -102.1233
25.8704 | -102.1071
25.8741 | -102.0888
25.8777 | -102.0731
25.8814 | -102.0609
25.885 | -102.0469
25.8887 | -102.0319
25.8924 | -102.0184
25.896 | -102.0049
25.8997 | -101.9924
25.9033 | -101.9804
25.907 | -101.9702
25.9107 | -101.9601
25.9143 | -101.9492
25.918 | -101.9378
25.9216 | -101.9278
25.9253 | -101.9192
25.9289 | -101.912
25.9326 | -101.9046
25.9363 | -101.8996
25.9399 | -101.8936
25.9436 | -101.8857
25.9472 | -101.8804
25.9509 | -101.8751
25.9545 | -101.8691
25.9582 | -101.8616
25.9619 | -101.8575
25.9655 | -101.8552
25.9692 | -101.8517
25.9728 | -101.8482
25.9765 | -101.845
25.9802 | -101.8435
25.9838 | -101.8439
25.9875 | -101.8458
25.9911 | -101.8484
25.9948 | -101.8507
25.9984 | -101.8526
26.0021 | -101.8542
26.0058 | -101.856
26.0094 | -101.8578
26.0131 | -101.8607
26.0167 | -101.8651
26.0204 | -101.8673
26.0241 | -101.8684
26.0277 | -101.8734
26.0314 | -101.8805
26.035 | -101.8871
26.0387 | -101.8912
26.0423 | -101.8919
26.046 | -101.8943
26.0497 | -101.897
26.0533 | -101.9005
26.057 | -101.9036
26.0606 | -101.9071
26.0643 | -101.9092
26.0679 | -101.9129
26.0716 | -101.9186
26.0753 | -101.9236
26.0789 | -101.9266
26.0826 | -101.9269
26.0862 | -101.9258
26.0899 | -101.9252
26.0936 | -101.9275
26.0972 | -101.9308
26.1009 | -101.9302
26.1045 | -101.9266
26.1082 | -101.9219
26.1118 | -101.9151
26.1155 | -101.9051
26.1192 | -101.8942
26.1228 | -101.8863
26.1265 | -101.8792
26.1301 | -101.8696
26.1338 | -101.8586
26.1375 | -101.8485
26.1411 | -101.8377
26.1448 | -101.8266
26.1484 | -101.8143
26.1521 | -101.8024
26.1557 | -101.7897
26.1594 | -101.7742
26.1631 | -101.7587
26.1667 | -101.7425
26.1704 | -101.7255
26.174 | -101.7104
26.1777 | -101.6938
26.1814 | -101.6753
26.185 | -101.6562
26.1887 | -101.6368
26.1923 | -101.6183
26.196 | -101.6012
26.1996 | -101.5815
26.2033 | -101.5606
26.207 | -101.5414
26.2106 | -101.521
26.2143 | -101.5014
26.2179 | -101.4814
26.2216 | -101.4649
26.2252 | -101.4477
26.2289 | -101.429
26.2326 | -101.41
26.2362 | -101.3879
26.2399 | -101.3697
26.2435 | -101.3542
26.2472 | -101.3379
26.2509 | -101.3205
26.2545 | -101.303
26.2582 | -101.2887
26.2618 | -101.2736
26.2655 | -101.2564
26.2691 | -101.24
26.2728 | -101.2266
26.2765 | -101.2161
26.2801 | -101.2048
26.2838 | -101.195
26.2874 | -101.1862
26.2911 | -101.1783
26.2948 | -101.1692
26.2984 | -101.1596
26.3021 | -101.1528
26.3057 | -101.146
26.3094 | -101.1386
26.313 | -101.134
26.3167 | -101.1268
26.3204 | -101.1228
26.324 | -101.12
26.3277 | -101.1163
26.3313 | -101.1101
26.335 | -101.1047
26.3387 | -101.1011
26.3423 | -101.099
26.346 | -101.0975
26.3496 | -101.096
26.3533 | -101.0941
26.3569 | -101.0914
26.3606 | -101.0893
26.3643 | -101.0872
26.3679 | -101.0851
26.3716 | -101.0827
26.3752 | -101.0804
26.3789 | -101.0788
26.3826 | -101.0782
26.3862 | -101.0783
26.3899 | -101.0784
26.3935 | -101.08
26.3972 | -101.0826
26.4008 | -101.0833
26.4045 | -101.0831
26.4082 | -101.0846
26.4118 | -101.0889
26.4155 | -101.0962
26.4191 | -101.1027
26.4228 | -101.1077
26.4264 | -101.1128
26.4301 | -101.1195
26.4338 | -101.1278
26.4374 | -101.1373
26.4411 | -101.1465
26.4447 | -101.1548
26.4484 | -101.1633
26.4521 | -101.1726
26.4557 | -101.1839
26.4594 | -101.1962
26.463 | -101.2087
26.4667 | -101.221
26.4703 | -101.2355
26.474 | -101.2507
26.4777 | -101.2667
26.4813 | -101.2846
26.485 | -101.3049
26.4886 | -101.3287
26.4923 | -101.3519
26.496 | -101.3725
26.4996 | -101.3909
26.5033 | -101.408
26.5069 | -101.4231
26.5106 | -101.4383
26.5142 | -101.4558
26.5179 | -101.4737
26.5216 | -101.4915
26.5252 | -101.5094
26.5289 | -101.5287
26.5325 | -101.5491
26.5362 | -101.5694
26.5399 | -101.5916
26.5435 | -101.6135
26.5472 | -101.6333
26.5508 | -101.6538
26.5545 | -101.6751
26.5581 | -101.6969
26.5618 | -101.7189
26.5655 | -101.7394
26.5691 | -101.759
26.5728 | -101.7789
26.5764 | -101.7982
26.5801 | -101.8161
26.5837 | -101.8362
26.5874 | -101.8563
26.5911 | -101.8772
26.5947 | -101.8958
26.5984 | -101.9138
26.602 | -101.9319
26.6057 | -101.9493
26.6094 | -101.9666
26.613 | -101.9825
26.6167 | -101.9966
26.6203 | -102.0095
26.624 | -102.0214
26.6276 | -102.036
26.6313 | -102.0494
26.635 | -102.0602
26.6386 | -102.0711
26.6423 | -102.0829
26.6459 | -102.0921
26.6496 | -102.0992
26.6533 | -102.1087
26.6569 | -102.1165
26.6606 | -102.125
26.6642 | -102.1329
26.6679 | -102.1429
26.6715 | -102.1537
26.6752 | -102.1617
26.6789 | -102.1709
26.6825 | -102.1815
26.6862 | -102.1878
26.6898 | -102.1963
26.6935 | -102.2033
26.6972 | -102.2108
26.7008 | -102.2215
26.7045 | -102.2335
26.7081 | -102.2426
26.7118 | -102.2499
26.7154 | -102.2591
26.7191 | -102.269
26.7228 | -102.2751
26.7264 | -102.2803
26.7301 | -102.2898
26.7337 | -102.2992
26.7374 | -102.31
26.741 | -102.3218
26.7447 | -102.3298
26.7484 | -102.3382
26.752 | -102.3487
26.7557 | -102.3587
26.7593 | -102.3691
26.763 | -102.3779
26.7667 | -102.3866
26.7703 | -102.393
26.774 | -102.3969
26.7776 | -102.4012
26.7813 | -102.4057
26.7849 | -102.4097
26.7886 | -102.4124
26.7923 | -102.4124
26.7959 | -102.4124
26.7996 | -102.4127
26.8032 | -102.4122
26.8069 | -102.412
26.8106 | -102.4079
26.8142 | -102.3982
26.8179 | -102.3942
26.8215 | -102.3935
26.8252 | -102.3887
26.8288 | -102.3816
26.8325 | -102.3753
26.8362 | -102.3688
26.8398 | -102.3603
26.8435 | -102.35
26.8471 | -102.3376
26.8508 | -102.3251
26.8545 | -102.3146
26.8581 | -102.305
26.8618 | -102.2954
26.8654 | -102.2819
26.8691 | -102.2683
26.8727 | -102.2553
26.8764 | -102.2426
26.8801 | -102.2316
26.8837 | -102.222
26.8874 | -102.2097
26.891 | -102.193
26.8947 | -102.177
26.8983 | -102.1618
26.902 | -102.1484
26.9057 | -102.1326
26.9093 | -102.1184
26.913 | -102.1042
26.9166 | -102.09
26.9203 | -102.0749
26.924 | -102.0599
26.9276 | -102.0446
26.9313 | -102.03
26.9349 | -102.0155
26.9386 | -102.0011
26.9422 | -101.9868
26.9459 | -101.9726
26.9496 | -101.9556
26.9532 | -101.9373
26.9569 | -101.9236
26.9605 | -101.909
26.9642 | -101.8926
26.9679 | -101.8762
26.9715 | -101.8604
26.9752 | -101.8462
26.9788 | -101.8341
26.9825 | -101.8207
26.9861 | -101.8059
26.9898 | -101.793
26.9935 | -101.7806
26.9971 | -101.7662
27.0008 | -101.7533
27.0044 | -101.7412
27.0081 | -101.7289
27.0118 | -101.7138
27.0154 | -101.7004
27.0191 | -101.6884
27.0227 | -101.6761
27.0264 | -101.6642
27.03 | -101.6522
27.0337 | -101.6391
27.0374 | -101.6273
27.041 | -101.6148
27.0447 | -101.6035
27.0483 | -101.5935
27.052 | -101.58
27.0556 | -101.5615
27.0593 | -101.5459
27.063 | -101.5347
27.0666 | -101.5255
27.0703 | -101.5152
27.0739 | -101.5045
27.0776 | -101.4938
27.0813 | -101.4824
27.0849 | -101.4698
27.0886 | -101.4595
27.0922 | -101.4511
27.0959 | -101.4411
27.0995 | -101.4297
27.1032 | -101.4198
27.1069 | -101.4122
27.1105 | -101.4033
27.1142 | -101.3936
27.1178 | -101.3846
27.1215 | -101.377
27.1252 | -101.3696
27.1288 | -101.3613
27.1325 | -101.3544
27.1361 | -101.3463
27.1398 | -101.3368
27.1434 | -101.3288
27.1471 | -101.3209
27.1508 | -101.3151
27.1544 | -101.3102
27.1581 | -101.3054
27.1617 | -101.3017
27.1654 | -101.2972
27.1691 | -101.2921
27.1727 | -101.2894
27.1764 | -101.2879
27.18 | -101.2854
27.1837 | -101.2824
27.1873 | -101.2806
27.191 | -101.2808
27.1947 | -101.2824
27.1983 | -101.2822
27.202 | -101.282
27.2056 | -101.2834
27.2093 | -101.2858
27.2129 | -101.2881
27.2166 | -101.2902
27.2203 | -101.2922
27.2239 | -101.2953
27.2276 | -101.2967
27.2312 | -101.2997
27.2349 | -101.3055
27.2386 | -101.3086
27.2422 | -101.3114
27.2459 | -101.3165
27.2495 | -101.3216
27.2532 | -101.3283
27.2568 | -101.3347
27.2605 | -101.3415
27.2642 | -101.3499
27.2678 | -101.359
27.2715 | -101.3701
27.2751 | -101.3788
27.2788 | -101.3875
27.2825 | -101.3921
27.2861 | -101.3985
27.2898 | -101.4075
27.2934 | -101.4148
27.2971 | -101.4247
27.3007 | -101.4341
27.3044 | -101.4421
27.3081 | -101.4503
27.3117 | -101.4585
27.3154 | -101.4659
27.319 | -101.4698
27.3227 | -101.4745
27.3264 | -101.4824
27.33 | -101.4903
27.3337 | -101.4997
27.3373 | -101.5093
27.341 | -101.5132
27.3446 | -101.5162
27.3483 | -101.5211
27.352 | -101.5249
27.3556 | -101.5266
27.3593 | -101.5282
27.3629 | -101.5282
27.3666 | -101.5277
27.3702 | -101.5256
27.3739 | -101.5216
27.3776 | -101.5185
27.3812 | -101.5189
27.3849 | -101.5167
27.3885 | -101.5144
27.3922 | -101.5105
27.3959 | -101.5049
27.3995 | -101.5017
27.4032 | -101.4984
27.4068 | -101.4929
27.4105 | -101.4869
27.4141 | -101.4814
27.4178 | -101.4772
27.4215 | -101.4713
27.4251 | -101.4674
27.4288 | -101.4627
27.4324 | -101.4554
27.4361 | -101.4523
27.4398 | -101.448
27.4434 | -101.4481
27.4471 | -101.4453
27.4507 | -101.4386
27.4544 | -101.4316
27.458 | -101.4276
27.4617 | -101.4231
27.4654 | -101.4164
27.469 | -101.4102
27.4727 | -101.4046
27.4763 | -101.3998
27.48 | -101.3949
27.4837 | -101.3867
27.4873 | -101.3799
27.491 | -101.3753
27.4946 | -101.369
27.4983 | -101.3625
27.5019 | -101.3577
27.5056 | -101.3535
27.5093 | -101.35
27.5129 | -101.3488
27.5166 | -101.3475
27.5202 | -101.3441
27.5239 | -101.3423
27.5276 | -101.3422
27.5312 | -101.3414
27.5349 | -101.3402
27.5385 | -101.3376
27.5422 | -101.3351
27.5458 | -101.334
27.5495 | -101.3322
27.5532 | -101.3305
27.5568 | -101.3294
27.5605 | -101.3286
27.5641 | -101.3264
27.5678 | -101.3275
27.5714 | -101.3289
27.5751 | -101.3282
27.5788 | -101.3273
27.5824 | -101.3272
27.5861 | -101.3271
27.5897 | -101.3269
27.5934 | -101.3274
27.5971 | -101.3274
27.6007 | -101.3265
27.6044 | -101.3277
27.608 | -101.3307
27.6117 | -101.3324
27.6153 | -101.336
27.619 | -101.337
27.6227 | -101.3341
27.6263 | -101.334
27.63 | -101.334
27.6336 | -101.3322
27.6373 | -101.3296
27.641 | -101.3252
27.6446 | -101.324
27.6483 | -101.3226
27.6519 | -101.3185
27.6556 | -101.3128
27.6592 | -101.306
27.6629 | -101.3032
27.6666 | -101.299
27.6702 | -101.2958
27.6739 | -101.2927
27.6775 | -101.2869
27.6812 | -101.2835
27.6849 | -101.2792
27.6885 | -101.274
27.6922 | -101.2684
27.6958 | -101.2616
27.6995 | -101.2555
27.7031 | -101.2495
27.7068 | -101.2457
27.7105 | -101.2413
27.7141 | -101.2337
27.7178 | -101.2261
27.7214 | -101.2196
27.7251 | -101.2143
27.7287 | -101.2081
27.7324 | -101.2031
27.7361 | -101.1984
27.7397 | -101.1938
27.7434 | -101.1894
27.747 | -101.1857
27.7507 | -101.1835
27.7544 | -101.1817
27.758 | -101.1812
27.7617 | -101.1808
27.7653 | -101.1792
27.769 | -101.1753
27.7726 | -101.1724
27.7763 | -101.1706
27.78 | -101.1712
27.7836 | -101.1741
27.7873 | -101.174
27.7909 | -101.1726
27.7946 | -101.1714
27.7983 | -101.171
27.8019 | -101.1685
27.8056 | -101.1646
27.8092 | -101.1633
27.8129 | -101.1605
27.8165 | -101.1579
27.8202 | -101.1597
27.8239 | -101.1606
27.8275 | -101.1612
27.8312 | -101.1638
27.8348 | -101.1637
27.8385 | -101.1629
27.8422 | -101.1661
27.8458 | -101.1686
27.8495 | -101.169
27.8531 | -101.1695
27.8568 | -101.1703
27.8604 | -101.1699
27.8641 | -101.1695
27.8678 | -101.1679
27.8714 | -101.1673
27.8751 | -101.1701
27.8787 | -101.173
27.8824 | -101.1747
27.886 | -101.1767
27.8897 | -101.1764
27.8934 | -101.1721
27.897 | -101.1666
27.9007 | -101.1578
27.9043 | -101.1469
27.908 | -101.1357
27.9117 | -101.1265
27.9153 | -101.1199
27.919 | -101.1109
27.9226 | -101.0965
27.9263 | -101.0811
27.9299 | -101.0646
27.9336 | -101.051
27.9373 | -101.0361
27.9409 | -101.0195
27.9446 | -101.0023
27.9482 | -100.9851
27.9519 | -100.9665
27.9556 | -100.9482
27.9592 | -100.9283
27.9629 | -100.9094
27.9665 | -100.8907
27.9702 | -100.8594
27.9738 | -100.833
27.9775 | -100.8067
27.9812 | -100.7817
27.9848 | -100.7585
27.9885 | -100.7355
27.9921 | -100.7123
27.9958 | -100.6886
27.9995 | -100.6632
28.0031 | -100.6371