## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 269 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 274
Hood | 02-92B/P | 214
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1012.6 | hPa |  | PRC PRO 01
Humidity | 57.1 | %/Hr
Internal temperature of calibration bench | 24.45 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1019 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 24.5 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2.01 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 20.03 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 6.02 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3072 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 10/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 17/07/2018
Technical Manager | Anthony HUE | Approval |  | 24/07/2018
CEO | Robert GLEMOT | Approval |  | 24/07/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.4294 | 57.6131
0.012209 | -35.9992 | 38.4456
0.018313 | -34.9399 | 28.2678
0.024417 | -34.4719 | 22.0013
0.030521 | -34.3574 | 18.2164
0.036626 | -34.2124 | 15.1744
0.04273 | -34.1166 | 12.8246
0.048834 | -34.0694 | 11.2966
0.054938 | -34.0046 | 10.1607
0.061043 | -33.945 | 9.0851
0.067147 | -33.9358 | 8.253
0.073251 | -33.9362 | 7.4918
0.079355 | -33.9519 | 6.8329
0.08546 | -33.9632 | 6.2383
0.091564 | -33.9574 | 5.7915
0.097668 | -33.9405 | 5.5335
0.10377 | -33.9204 | 5.2811
0.10988 | -33.9042 | 5.0004
0.11598 | -33.8951 | 4.7075
0.12209 | -33.8962 | 4.4264
0.12819 | -33.9051 | 4.1487
0.13429 | -33.9151 | 3.8686
0.1404 | -33.918 | 3.5974
0.1465 | -33.9197 | 3.3448
0.15261 | -33.917 | 3.1268
0.15871 | -33.9107 | 2.9627
0.16482 | -33.903 | 2.8546
0.17092 | -33.9008 | 2.8088
0.17702 | -33.8986 | 2.7325
0.18313 | -33.8977 | 2.6523
0.18923 | -33.8972 | 2.5447
0.19534 | -33.8959 | 2.4213
0.20144 | -33.8945 | 2.2835
0.20754 | -33.8917 | 2.1641
0.21365 | -33.8908 | 2.0508
0.21975 | -33.8903 | 1.9529
0.22586 | -33.8913 | 1.865
0.23196 | -33.8938 | 1.7918
0.23807 | -33.8976 | 1.728
0.24417 | -33.9014 | 1.6695
0.25027 | -33.905 | 1.6069
0.25638 | -33.9071 | 1.5451
0.26248 | -33.9087 | 1.479
0.26859 | -33.9094 | 1.4143
0.27469 | -33.9096 | 1.3441
0.2808 | -33.9106 | 1.2718
0.2869 | -33.9113 | 1.2021
0.293 | -33.9122 | 1.1358
0.29911 | -33.9132 | 1.0764
0.30521 | -33.9142 | 1.0178
0.31132 | -33.9151 | 0.97846
0.31742 | -33.9163 | 0.93519
0.32353 | -33.9167 | 0.89934
0.32963 | -33.9176 | 0.86919
0.33573 | -33.9186 | 0.84732
0.34184 | -33.9198 | 0.82469
0.34794 | -33.9217 | 0.80019
0.35405 | -33.9237 | 0.77536
0.36015 | -33.926 | 0.74928
0.36626 | -33.9279 | 0.72112
0.37236 | -33.9298 | 0.68438
0.37846 | -33.9318 | 0.64044
0.38457 | -33.9326 | 0.58436
0.39067 | -33.9335 | 0.5382
0.39678 | -33.9341 | 0.48716
0.40288 | -33.9345 | 0.4341
0.40899 | -33.9344 | 0.38324
0.41509 | -33.9337 | 0.33338
0.42119 | -33.9337 | 0.29116
0.4273 | -33.9336 | 0.24597
0.4334 | -33.9335 | 0.20109
0.43951 | -33.9337 | 0.1587
0.44561 | -33.9344 | 0.11981
0.45172 | -33.9358 | 0.088485
0.45782 | -33.9366 | 0.051947
0.46392 | -33.9377 | 0.01628
0.47003 | -33.9387 | -0.015796
0.47613 | -33.9395 | -0.049725
0.48224 | -33.94 | -0.084286
0.48834 | -33.9391 | -0.11927
0.49445 | -33.9389 | -0.15448
0.50055 | -33.9384 | -0.18669
0.50665 | -33.9377 | -0.22097
0.51276 | -33.9368 | -0.25493
0.51886 | -33.9358 | -0.28882
0.52497 | -33.9347 | -0.31321
0.53107 | -33.934 | -0.34264
0.53717 | -33.9333 | -0.37078
0.54328 | -33.9325 | -0.39379
0.54938 | -33.9324 | -0.4164
0.55549 | -33.9324 | -0.43744
0.56159 | -33.9335 | -0.44934
0.5677 | -33.9338 | -0.46431
0.5738 | -33.9348 | -0.48082
0.5799 | -33.9356 | -0.49599
0.58601 | -33.9366 | -0.50645
0.59211 | -33.9382 | -0.51345
0.59822 | -33.9395 | -0.52345
0.60432 | -33.9409 | -0.53592
0.61043 | -33.9421 | -0.54445
0.61653 | -33.9433 | -0.55516
0.62263 | -33.9448 | -0.56501
0.62874 | -33.9454 | -0.57607
0.63484 | -33.9463 | -0.58584
0.64095 | -33.9467 | -0.59641
0.64705 | -33.9471 | -0.60662
0.65316 | -33.9473 | -0.61941
0.65926 | -33.9474 | -0.63526
0.66536 | -33.9463 | -0.65267
0.67147 | -33.9459 | -0.66597
0.67757 | -33.9456 | -0.68041
0.68368 | -33.9449 | -0.69925
0.68978 | -33.9441 | -0.71634
0.69589 | -33.9432 | -0.73525
0.70199 | -33.9423 | -0.75771
0.70809 | -33.9415 | -0.77671
0.7142 | -33.9408 | -0.79707
0.7203 | -33.9397 | -0.81621
0.72641 | -33.9391 | -0.83856
0.73251 | -33.9384 | -0.85856
0.73862 | -33.9386 | -0.88037
0.74472 | -33.9383 | -0.90046
0.75082 | -33.938 | -0.92257
0.75693 | -33.9383 | -0.94432
0.76303 | -33.9388 | -0.96635
0.76914 | -33.9406 | -0.98704
0.77524 | -33.9415 | -1.0069
0.78135 | -33.9426 | -1.0271
0.78745 | -33.9439 | -1.0472
0.79355 | -33.9451 | -1.0672
0.79966 | -33.9467 | -1.0868
0.80576 | -33.9489 | -1.1067
0.81187 | -33.9504 | -1.1283
0.81797 | -33.9523 | -1.1497
0.82408 | -33.9539 | -1.1704
0.83018 | -33.9555 | -1.191
0.83628 | -33.9573 | -1.2097
0.84239 | -33.9584 | -1.23
0.84849 | -33.9598 | -1.2517
0.8546 | -33.9611 | -1.2715
0.8607 | -33.9623 | -1.2922
0.86681 | -33.9633 | -1.3112
0.87291 | -33.9642 | -1.33
0.87901 | -33.9647 | -1.3478
0.88512 | -33.9653 | -1.3659
0.89122 | -33.966 | -1.3841
0.89733 | -33.9664 | -1.3998
0.90343 | -33.9668 | -1.4174
0.90953 | -33.9669 | -1.4317
0.91564 | -33.9669 | -1.4475
0.92174 | -33.9671 | -1.4625
0.92785 | -33.9672 | -1.477
0.93395 | -33.9672 | -1.4909
0.94006 | -33.9674 | -1.5038
0.94616 | -33.9671 | -1.5157
0.95226 | -33.9672 | -1.5284
0.95837 | -33.9672 | -1.5422
0.96447 | -33.9674 | -1.555
0.97058 | -33.9672 | -1.5684
0.97668 | -33.967 | -1.5817
0.98279 | -33.9667 | -1.5944
0.98889 | -33.9666 | -1.6051
0.99499 | -33.9665 | -1.6185
1.0011 | -33.9664 | -1.6292
1.0072 | -33.966 | -1.6409
1.0133 | -33.9657 | -1.6516
1.0194 | -33.9653 | -1.6646
1.0255 | -33.965 | -1.6782
1.0316 | -33.9644 | -1.6901
1.0377 | -33.9641 | -1.7037
1.0438 | -33.9635 | -1.7163
1.0499 | -33.9628 | -1.7314
1.056 | -33.9621 | -1.7456
1.0621 | -33.9616 | -1.7598
1.0682 | -33.9611 | -1.7723
1.0743 | -33.9605 | -1.7876
1.0805 | -33.9601 | -1.8019
1.0866 | -33.9592 | -1.817
1.0927 | -33.9584 | -1.8315
1.0988 | -33.9578 | -1.8463
1.1049 | -33.9571 | -1.8605
1.111 | -33.9564 | -1.8759
1.1171 | -33.9557 | -1.8905
1.1232 | -33.9553 | -1.9057
1.1293 | -33.9549 | -1.9198
1.1354 | -33.9547 | -1.9349
1.1415 | -33.9543 | -1.9504
1.1476 | -33.9538 | -1.9659
1.1537 | -33.9533 | -1.9806
1.1598 | -33.9534 | -1.9955
1.1659 | -33.9532 | -2.009
1.172 | -33.953 | -2.023
1.1781 | -33.9528 | -2.0371
1.1842 | -33.9527 | -2.0513
1.1903 | -33.953 | -2.0625
1.1964 | -33.9529 | -2.0761
1.2025 | -33.953 | -2.0901
1.2086 | -33.9533 | -2.1027
1.2147 | -33.9533 | -2.1154
1.2209 | -33.9537 | -2.1267
1.227 | -33.9544 | -2.1377
1.2331 | -33.9547 | -2.1498
1.2392 | -33.9551 | -2.1617
1.2453 | -33.9555 | -2.1729
1.2514 | -33.9559 | -2.183
1.2575 | -33.9564 | -2.1941
1.2636 | -33.9572 | -2.2032
1.2697 | -33.9577 | -2.2127
1.2758 | -33.9584 | -2.2238
1.2819 | -33.9591 | -2.2339
1.288 | -33.9597 | -2.2448
1.2941 | -33.9604 | -2.2534
1.3002 | -33.9611 | -2.2616
1.3063 | -33.9616 | -2.2708
1.3124 | -33.9622 | -2.28
1.3185 | -33.9629 | -2.2892
1.3246 | -33.9635 | -2.2984
1.3307 | -33.9643 | -2.3071
1.3368 | -33.965 | -2.3159
1.3429 | -33.9656 | -2.325
1.349 | -33.9661 | -2.3344
1.3551 | -33.9667 | -2.3438
1.3613 | -33.9673 | -2.3522
1.3674 | -33.9679 | -2.3623
1.3735 | -33.9683 | -2.3721
1.3796 | -33.9687 | -2.3813
1.3857 | -33.969 | -2.3907
1.3918 | -33.9695 | -2.401
1.3979 | -33.9701 | -2.4103
1.404 | -33.9705 | -2.4217
1.4101 | -33.9709 | -2.4322
1.4162 | -33.9713 | -2.4428
1.4223 | -33.9717 | -2.4542
1.4284 | -33.9722 | -2.4644
1.4345 | -33.9724 | -2.476
1.4406 | -33.9727 | -2.4895
1.4467 | -33.9732 | -2.5012
1.4528 | -33.9736 | -2.5127
1.4589 | -33.974 | -2.5246
1.465 | -33.9743 | -2.5355
1.4711 | -33.9747 | -2.547
1.4772 | -33.9752 | -2.5603
1.4833 | -33.9757 | -2.5733
1.4894 | -33.9762 | -2.5856
1.4955 | -33.9767 | -2.5976
1.5016 | -33.9772 | -2.6102
1.5078 | -33.9776 | -2.624
1.5139 | -33.978 | -2.6371
1.52 | -33.9784 | -2.6497
1.5261 | -33.9787 | -2.6624
1.5322 | -33.979 | -2.6763
1.5383 | -33.9794 | -2.6889
1.5444 | -33.9799 | -2.7016
1.5505 | -33.9804 | -2.7151
1.5566 | -33.9809 | -2.7281
1.5627 | -33.9813 | -2.7419
1.5688 | -33.9819 | -2.7554
1.5749 | -33.9822 | -2.7681
1.581 | -33.9825 | -2.7815
1.5871 | -33.9829 | -2.7952
1.5932 | -33.9832 | -2.8081
1.5993 | -33.9836 | -2.8222
1.6054 | -33.984 | -2.8355
1.6115 | -33.9844 | -2.8495
1.6176 | -33.9846 | -2.8621
1.6237 | -33.9849 | -2.8757
1.6298 | -33.9851 | -2.8883
1.6359 | -33.9855 | -2.9011
1.642 | -33.9859 | -2.9132
1.6482 | -33.986 | -2.9262
1.6543 | -33.9864 | -2.9399
1.6604 | -33.9867 | -2.9519
1.6665 | -33.9867 | -2.9644
1.6726 | -33.987 | -2.977
1.6787 | -33.9873 | -2.9896
1.6848 | -33.9872 | -3.0022
1.6909 | -33.9875 | -3.0149
1.697 | -33.9877 | -3.0272
1.7031 | -33.9877 | -3.0398
1.7092 | -33.9877 | -3.0521
1.7153 | -33.9878 | -3.0641
1.7214 | -33.9877 | -3.0762
1.7275 | -33.9876 | -3.0888
1.7336 | -33.9876 | -3.1017
1.7397 | -33.9877 | -3.1141
1.7458 | -33.9877 | -3.1263
1.7519 | -33.9878 | -3.1384
1.758 | -33.9878 | -3.1503
1.7641 | -33.988 | -3.1619
1.7702 | -33.9879 | -3.1748
1.7763 | -33.9879 | -3.1863
1.7824 | -33.9879 | -3.1987
1.7885 | -33.9879 | -3.2112
1.7947 | -33.988 | -3.2231
1.8008 | -33.988 | -3.2344
1.8069 | -33.988 | -3.246
1.813 | -33.9881 | -3.2581
1.8191 | -33.9881 | -3.2707
1.8252 | -33.9882 | -3.2833
1.8313 | -33.9883 | -3.2955
1.8374 | -33.9885 | -3.3073
1.8435 | -33.9885 | -3.3191
1.8496 | -33.9886 | -3.332
1.8557 | -33.9888 | -3.3446
1.8618 | -33.9891 | -3.3567
1.8679 | -33.9892 | -3.3686
1.874 | -33.9893 | -3.3805
1.8801 | -33.9895 | -3.3927
1.8862 | -33.9898 | -3.4046
1.8923 | -33.9899 | -3.4166
1.8984 | -33.99 | -3.4285
1.9045 | -33.9901 | -3.4399
1.9106 | -33.9902 | -3.4519
1.9167 | -33.9902 | -3.4639
1.9228 | -33.9904 | -3.4757
1.9289 | -33.9905 | -3.4873
1.9351 | -33.9907 | -3.4981
1.9412 | -33.9909 | -3.5092
1.9473 | -33.9912 | -3.5211
1.9534 | -33.9914 | -3.5324
1.9595 | -33.9916 | -3.544
1.9656 | -33.9918 | -3.5561
1.9717 | -33.9921 | -3.5677
1.9778 | -33.9924 | -3.5789
1.9839 | -33.9926 | -3.5902
1.99 | -33.9928 | -3.6012
1.9961 | -33.9929 | -3.6131
2.0022 | -33.9931 | -3.6244
2.0083 | -33.9932 | -3.6355
2.0144 | -33.9935 | -3.6468
2.0205 | -33.9938 | -3.6584
2.0266 | -33.9941 | -3.6697
2.0327 | -33.9943 | -3.6811
2.0388 | -33.9947 | -3.6926
2.0449 | -33.9949 | -3.7039
2.051 | -33.9952 | -3.7152
2.0571 | -33.9955 | -3.7264
2.0632 | -33.9957 | -3.7378
2.0693 | -33.9959 | -3.7498
2.0754 | -33.9961 | -3.7611
2.0816 | -33.9964 | -3.7731
2.0877 | -33.9966 | -3.7846
2.0938 | -33.9969 | -3.7966
2.0999 | -33.9971 | -3.8082
2.106 | -33.9973 | -3.8198
2.1121 | -33.9975 | -3.8317
2.1182 | -33.9977 | -3.8425
2.1243 | -33.9979 | -3.8532
2.1304 | -33.9981 | -3.8643
2.1365 | -33.9984 | -3.876
2.1426 | -33.9986 | -3.888
2.1487 | -33.9989 | -3.8992
2.1548 | -33.9993 | -3.911
2.1609 | -33.9997 | -3.9224
2.167 | -34 | -3.9337
2.1731 | -34.0003 | -3.9456
2.1792 | -34.0005 | -3.9568
2.1853 | -34.0009 | -3.9687
2.1914 | -34.0011 | -3.98
2.1975 | -34.0015 | -3.9915
2.2036 | -34.0017 | -4.0036
2.2097 | -34.0019 | -4.0151
2.2158 | -34.0023 | -4.0257
2.222 | -34.0025 | -4.0362
2.2281 | -34.0028 | -4.0476
2.2342 | -34.003 | -4.059
2.2403 | -34.0032 | -4.0696
2.2464 | -34.0035 | -4.0805
2.2525 | -34.0037 | -4.0919
2.2586 | -34.004 | -4.1026
2.2647 | -34.0043 | -4.1131
2.2708 | -34.0046 | -4.1244
2.2769 | -34.0049 | -4.1351
2.283 | -34.0051 | -4.1463
2.2891 | -34.0053 | -4.1572
2.2952 | -34.0054 | -4.1679
2.3013 | -34.0056 | -4.1787
2.3074 | -34.0059 | -4.1892
2.3135 | -34.0061 | -4.2007
2.3196 | -34.0063 | -4.2109
2.3257 | -34.0064 | -4.2216
2.3318 | -34.0066 | -4.2326
2.3379 | -34.0067 | -4.2434
2.344 | -34.0069 | -4.2548
2.3501 | -34.0071 | -4.2656
2.3562 | -34.0072 | -4.276
2.3623 | -34.0073 | -4.2865
2.3685 | -34.0074 | -4.2971
2.3746 | -34.0076 | -4.3083
2.3807 | -34.0078 | -4.3189
2.3868 | -34.0079 | -4.3296
2.3929 | -34.0081 | -4.3403
2.399 | -34.0082 | -4.3508
2.4051 | -34.0084 | -4.3614
2.4112 | -34.0086 | -4.3721
2.4173 | -34.0087 | -4.3828
2.4234 | -34.0088 | -4.3936
2.4295 | -34.009 | -4.4045
2.4356 | -34.0091 | -4.4154
2.4417 | -34.0092 | -4.4265
2.4478 | -34.0093 | -4.4374
2.4539 | -34.0095 | -4.4477
2.46 | -34.0096 | -4.4592
2.4661 | -34.0097 | -4.4703
2.4722 | -34.0099 | -4.4806
2.4783 | -34.01 | -4.4911
2.4844 | -34.01 | -4.5022
2.4905 | -34.0101 | -4.5132
2.4966 | -34.0103 | -4.5239
2.5027 | -34.0104 | -4.5345
2.5089 | -34.0105 | -4.5454
2.515 | -34.0105 | -4.5562
2.5211 | -34.0106 | -4.567
2.5272 | -34.0107 | -4.578
2.5333 | -34.0109 | -4.589
2.5394 | -34.011 | -4.5998
2.5455 | -34.0111 | -4.6112
2.5516 | -34.0113 | -4.6216
2.5577 | -34.0115 | -4.6331
2.5638 | -34.0116 | -4.6448
2.5699 | -34.0117 | -4.6562
2.576 | -34.0117 | -4.6679
2.5821 | -34.0118 | -4.6792
2.5882 | -34.012 | -4.6904
2.5943 | -34.0121 | -4.7013
2.6004 | -34.0123 | -4.7124
2.6065 | -34.0124 | -4.7228
2.6126 | -34.0125 | -4.7334
2.6187 | -34.0126 | -4.744
2.6248 | -34.0128 | -4.7549
2.6309 | -34.0129 | -4.7663
2.637 | -34.0132 | -4.778
2.6431 | -34.0134 | -4.7885
2.6492 | -34.0135 | -4.7999
2.6554 | -34.0137 | -4.8108
2.6615 | -34.0138 | -4.8219
2.6676 | -34.014 | -4.8334
2.6737 | -34.0143 | -4.8446
2.6798 | -34.0146 | -4.8558
2.6859 | -34.0147 | -4.8672
2.692 | -34.0149 | -4.8781
2.6981 | -34.0151 | -4.8892
2.7042 | -34.0153 | -4.9004
2.7103 | -34.0155 | -4.9111
2.7164 | -34.0158 | -4.9221
2.7225 | -34.0161 | -4.9335
2.7286 | -34.0163 | -4.9446
2.7347 | -34.0165 | -4.9552
2.7408 | -34.0168 | -4.9654
2.7469 | -34.017 | -4.9767
2.753 | -34.0173 | -4.9874
2.7591 | -34.0176 | -4.9987
2.7652 | -34.0178 | -5.0103
2.7713 | -34.018 | -5.0211
2.7774 | -34.0183 | -5.032
2.7835 | -34.0184 | -5.043
2.7896 | -34.0187 | -5.0536
2.7958 | -34.0189 | -5.0644
2.8019 | -34.0192 | -5.0749
2.808 | -34.0195 | -5.0857
2.8141 | -34.0198 | -5.0969
2.8202 | -34.0201 | -5.1079
2.8263 | -34.0204 | -5.1188
2.8324 | -34.0207 | -5.1302
2.8385 | -34.0211 | -5.1411
2.8446 | -34.0215 | -5.1518
2.8507 | -34.0218 | -5.1626
2.8568 | -34.0222 | -5.1732
2.8629 | -34.0224 | -5.1838
2.869 | -34.0228 | -5.1949
2.8751 | -34.0231 | -5.2056
2.8812 | -34.0235 | -5.2168
2.8873 | -34.0237 | -5.2277
2.8934 | -34.0241 | -5.2387
2.8995 | -34.0244 | -5.2498
2.9056 | -34.0246 | -5.2609
2.9117 | -34.025 | -5.2719
2.9178 | -34.0253 | -5.2826
2.9239 | -34.0257 | -5.2941
2.93 | -34.0261 | -5.3055
2.9361 | -34.0264 | -5.3163
2.9423 | -34.0267 | -5.3269
2.9484 | -34.0271 | -5.338
2.9545 | -34.0275 | -5.3489
2.9606 | -34.0278 | -5.3602
2.9667 | -34.0282 | -5.3713
2.9728 | -34.0285 | -5.3826
2.9789 | -34.0289 | -5.3936
2.985 | -34.0293 | -5.4047
2.9911 | -34.0297 | -5.4159
2.9972 | -34.0301 | -5.4269
3.0033 | -34.0305 | -5.4383
3.0094 | -34.0308 | -5.4493
3.0155 | -34.0312 | -5.4602
3.0216 | -34.0316 | -5.4714
3.0277 | -34.0319 | -5.483
3.0338 | -34.0323 | -5.4944
3.0399 | -34.0327 | -5.5056
3.046 | -34.0329 | -5.517
3.0521 | -34.0333 | -5.528
3.0582 | -34.0335 | -5.5399
3.0643 | -34.0339 | -5.5512
3.0704 | -34.0341 | -5.563
3.0765 | -34.0344 | -5.5739
3.0827 | -34.0347 | -5.585
3.0888 | -34.035 | -5.5962
3.0949 | -34.0354 | -5.6074
3.101 | -34.0357 | -5.6187
3.1071 | -34.036 | -5.6303
3.1132 | -34.0364 | -5.642
3.1193 | -34.0367 | -5.6533
3.1254 | -34.037 | -5.6646
3.1315 | -34.0374 | -5.6763
3.1376 | -34.0376 | -5.6876
3.1437 | -34.038 | -5.6991
3.1498 | -34.0383 | -5.711
3.1559 | -34.0386 | -5.7223
3.162 | -34.0388 | -5.7343
3.1681 | -34.039 | -5.7461
3.1742 | -34.0393 | -5.7575
3.1803 | -34.0396 | -5.7691
3.1864 | -34.0399 | -5.7802
3.1925 | -34.0402 | -5.7916
3.1986 | -34.0405 | -5.8034
3.2047 | -34.0407 | -5.8153
3.2108 | -34.0409 | -5.8264
3.2169 | -34.0412 | -5.838
3.223 | -34.0414 | -5.8495
3.2292 | -34.0416 | -5.8612
3.2353 | -34.0418 | -5.872
3.2414 | -34.042 | -5.8834
3.2475 | -34.0422 | -5.8947
3.2536 | -34.0425 | -5.906
3.2597 | -34.0427 | -5.917
3.2658 | -34.0429 | -5.9278
3.2719 | -34.0431 | -5.9389
3.278 | -34.0434 | -5.9504
3.2841 | -34.0436 | -5.9615
3.2902 | -34.0438 | -5.9728
3.2963 | -34.044 | -5.9842
3.3024 | -34.0441 | -5.9953
3.3085 | -34.0444 | -6.0068
3.3146 | -34.0445 | -6.0179
3.3207 | -34.0448 | -6.0293
3.3268 | -34.045 | -6.0407
3.3329 | -34.0451 | -6.0514
3.339 | -34.0453 | -6.0622
3.3451 | -34.0455 | -6.0734
3.3512 | -34.0456 | -6.0841
3.3573 | -34.0458 | -6.0953
3.3634 | -34.046 | -6.1064
3.3696 | -34.0462 | -6.1177
3.3757 | -34.0463 | -6.1285
3.3818 | -34.0465 | -6.1394
3.3879 | -34.0467 | -6.1504
3.394 | -34.0468 | -6.1618
3.4001 | -34.047 | -6.1727
3.4062 | -34.0471 | -6.1836
3.4123 | -34.0472 | -6.1942
3.4184 | -34.0474 | -6.2055
3.4245 | -34.0476 | -6.2163
3.4306 | -34.0477 | -6.227
3.4367 | -34.0479 | -6.2375
3.4428 | -34.0481 | -6.248
3.4489 | -34.0482 | -6.2589
3.455 | -34.0483 | -6.2698
3.4611 | -34.0485 | -6.2811
3.4672 | -34.0486 | -6.2922
3.4733 | -34.0488 | -6.3033
3.4794 | -34.049 | -6.3141
3.4855 | -34.0491 | -6.3246
3.4916 | -34.0493 | -6.335
3.4977 | -34.0494 | -6.3464
3.5038 | -34.0496 | -6.3573
3.5099 | -34.0497 | -6.3676
3.5161 | -34.0499 | -6.3781
3.5222 | -34.05 | -6.3886
3.5283 | -34.0501 | -6.3994
3.5344 | -34.0502 | -6.4103
3.5405 | -34.0504 | -6.4209
3.5466 | -34.0506 | -6.4318
3.5527 | -34.0507 | -6.4428
3.5588 | -34.0508 | -6.4535
3.5649 | -34.0508 | -6.4643
3.571 | -34.051 | -6.4749
3.5771 | -34.0511 | -6.4861
3.5832 | -34.0513 | -6.496
3.5893 | -34.0515 | -6.5067
3.5954 | -34.0516 | -6.5176
3.6015 | -34.0518 | -6.5285
3.6076 | -34.052 | -6.5391
3.6137 | -34.0522 | -6.5493
3.6198 | -34.0524 | -6.5597
3.6259 | -34.0525 | -6.5706
3.632 | -34.0527 | -6.5813
3.6381 | -34.0529 | -6.5921
3.6442 | -34.0531 | -6.6027
3.6503 | -34.0533 | -6.613
3.6565 | -34.0535 | -6.6238
3.6626 | -34.0536 | -6.6344
3.6687 | -34.0539 | -6.6452
3.6748 | -34.0541 | -6.6556
3.6809 | -34.0544 | -6.6664
3.687 | -34.0545 | -6.6774
3.6931 | -34.0548 | -6.6877
3.6992 | -34.055 | -6.6984
3.7053 | -34.0552 | -6.7089
3.7114 | -34.0554 | -6.7197
3.7175 | -34.0556 | -6.7305
3.7236 | -34.0559 | -6.7406
3.7297 | -34.0561 | -6.7514
3.7358 | -34.0563 | -6.7624
3.7419 | -34.0565 | -6.7732
3.748 | -34.0567 | -6.7838
3.7541 | -34.0569 | -6.7945
3.7602 | -34.0572 | -6.8051
3.7663 | -34.0574 | -6.8156
3.7724 | -34.0577 | -6.8264
3.7785 | -34.058 | -6.8369
3.7846 | -34.0581 | -6.8475
3.7907 | -34.0584 | -6.8584
3.7969 | -34.0587 | -6.8692
3.803 | -34.0589 | -6.8794
3.8091 | -34.0592 | -6.89
3.8152 | -34.0594 | -6.9005
3.8213 | -34.0597 | -6.9114
3.8274 | -34.0599 | -6.9219
3.8335 | -34.0603 | -6.9326
3.8396 | -34.0606 | -6.9434
3.8457 | -34.0609 | -6.9542
3.8518 | -34.0612 | -6.9646
3.8579 | -34.0615 | -6.9753
3.864 | -34.0618 | -6.9861
3.8701 | -34.0621 | -6.9971
3.8762 | -34.0624 | -7.0076
3.8823 | -34.0626 | -7.0185
3.8884 | -34.0629 | -7.0291
3.8945 | -34.0633 | -7.0401
3.9006 | -34.0635 | -7.051
3.9067 | -34.0638 | -7.0621
3.9128 | -34.0641 | -7.0729
3.9189 | -34.0644 | -7.0834
3.925 | -34.0646 | -7.0942
3.9311 | -34.0649 | -7.1045
3.9372 | -34.0652 | -7.1154
3.9434 | -34.0655 | -7.126
3.9495 | -34.0658 | -7.1369
3.9556 | -34.0661 | -7.1474
3.9617 | -34.0664 | -7.1578
3.9678 | -34.0668 | -7.1682
3.9739 | -34.067 | -7.1788
3.98 | -34.0673 | -7.1896
3.9861 | -34.0675 | -7.2003
3.9922 | -34.0678 | -7.211
3.9983 | -34.0681 | -7.2214
4.0044 | -34.0683 | -7.2317
4.0105 | -34.0686 | -7.2423
4.0166 | -34.0689 | -7.2531
4.0227 | -34.0692 | -7.2637
4.0288 | -34.0695 | -7.2747
4.0349 | -34.0697 | -7.2852
4.041 | -34.0699 | -7.2958
4.0471 | -34.0702 | -7.3066
4.0532 | -34.0704 | -7.3169
4.0593 | -34.0708 | -7.3278
4.0654 | -34.071 | -7.3382
4.0715 | -34.0713 | -7.3493
4.0776 | -34.0716 | -7.3599
4.0838 | -34.0719 | -7.371
4.0899 | -34.0722 | -7.3814
4.096 | -34.0724 | -7.392
4.1021 | -34.0727 | -7.4024
4.1082 | -34.0729 | -7.4129
4.1143 | -34.0732 | -7.4231
4.1204 | -34.0734 | -7.4335
4.1265 | -34.0737 | -7.4439
4.1326 | -34.074 | -7.4543
4.1387 | -34.0743 | -7.4648
4.1448 | -34.0745 | -7.4754
4.1509 | -34.0747 | -7.4861
4.157 | -34.075 | -7.4971
4.1631 | -34.0752 | -7.5079
4.1692 | -34.0755 | -7.5183
4.1753 | -34.0757 | -7.5289
4.1814 | -34.0759 | -7.5391
4.1875 | -34.0762 | -7.55
4.1936 | -34.0764 | -7.5606
4.1997 | -34.0767 | -7.5708
4.2058 | -34.0769 | -7.5816
4.2119 | -34.0771 | -7.592
4.218 | -34.0773 | -7.6023
4.2241 | -34.0775 | -7.6127
4.2303 | -34.0778 | -7.623
4.2364 | -34.078 | -7.6336
4.2425 | -34.0782 | -7.6442
4.2486 | -34.0784 | -7.6548
4.2547 | -34.0785 | -7.6652
4.2608 | -34.0787 | -7.6758
4.2669 | -34.0789 | -7.686
4.273 | -34.0791 | -7.6967
4.2791 | -34.0793 | -7.7072
4.2852 | -34.0796 | -7.7179
4.2913 | -34.0797 | -7.7281
4.2974 | -34.0799 | -7.7389
4.3035 | -34.0801 | -7.7497
4.3096 | -34.0803 | -7.7603
4.3157 | -34.0805 | -7.7706
4.3218 | -34.0807 | -7.7809
4.3279 | -34.0809 | -7.7914
4.334 | -34.0811 | -7.8017
4.3401 | -34.0813 | -7.8119
4.3462 | -34.0815 | -7.8225
4.3523 | -34.0817 | -7.8325
4.3584 | -34.0819 | -7.8426
4.3645 | -34.0821 | -7.853
4.3707 | -34.0823 | -7.8632
4.3768 | -34.0825 | -7.8735
4.3829 | -34.0827 | -7.8843
4.389 | -34.0829 | -7.8946
4.3951 | -34.083 | -7.9048
4.4012 | -34.0832 | -7.9155
4.4073 | -34.0834 | -7.9257
4.4134 | -34.0835 | -7.936
4.4195 | -34.0838 | -7.9463
4.4256 | -34.084 | -7.9565
4.4317 | -34.0842 | -7.9668
4.4378 | -34.0844 | -7.9768
4.4439 | -34.0846 | -7.9869
4.45 | -34.0847 | -7.9972
4.4561 | -34.0848 | -8.0075
4.4622 | -34.085 | -8.0178
4.4683 | -34.0852 | -8.0281
4.4744 | -34.0854 | -8.0388
4.4805 | -34.0856 | -8.0487
4.4866 | -34.0858 | -8.0591
4.4927 | -34.086 | -8.0696
4.4988 | -34.0862 | -8.0798
4.5049 | -34.0864 | -8.0897
4.511 | -34.0866 | -8.0997
4.5172 | -34.0868 | -8.11
4.5233 | -34.087 | -8.1206
4.5294 | -34.0872 | -8.1308
4.5355 | -34.0874 | -8.141
4.5416 | -34.0876 | -8.1513
4.5477 | -34.0878 | -8.1616
4.5538 | -34.0879 | -8.1715
4.5599 | -34.0882 | -8.1817
4.566 | -34.0884 | -8.192
4.5721 | -34.0886 | -8.2023
4.5782 | -34.0888 | -8.2127
4.5843 | -34.089 | -8.2231
4.5904 | -34.0893 | -8.2334
4.5965 | -34.0895 | -8.2435
4.6026 | -34.0897 | -8.2541
4.6087 | -34.0899 | -8.2643
4.6148 | -34.09 | -8.2749
4.6209 | -34.0903 | -8.2851
4.627 | -34.0905 | -8.2956
4.6331 | -34.0907 | -8.3057
4.6392 | -34.0909 | -8.3161
4.6453 | -34.0911 | -8.3263
4.6514 | -34.0913 | -8.3366
4.6576 | -34.0915 | -8.3473
4.6637 | -34.0917 | -8.3572
4.6698 | -34.0919 | -8.3673
4.6759 | -34.0921 | -8.378
4.682 | -34.0923 | -8.3881
4.6881 | -34.0924 | -8.3983
4.6942 | -34.0926 | -8.4085
4.7003 | -34.0929 | -8.4189
4.7064 | -34.0931 | -8.429
4.7125 | -34.0933 | -8.4396
4.7186 | -34.0935 | -8.4497
4.7247 | -34.0938 | -8.4601
4.7308 | -34.094 | -8.4703
4.7369 | -34.0942 | -8.4806
4.743 | -34.0944 | -8.4913
4.7491 | -34.0946 | -8.5019
4.7552 | -34.0948 | -8.5123
4.7613 | -34.095 | -8.5227
4.7674 | -34.0952 | -8.5329
4.7735 | -34.0954 | -8.5436
4.7796 | -34.0957 | -8.5541
4.7857 | -34.0959 | -8.5645
4.7918 | -34.0962 | -8.5747
4.7979 | -34.0964 | -8.5849
4.8041 | -34.0967 | -8.5953
4.8102 | -34.0969 | -8.6053
4.8163 | -34.0971 | -8.6156
4.8224 | -34.0974 | -8.626
4.8285 | -34.0976 | -8.636
4.8346 | -34.0979 | -8.6462
4.8407 | -34.0981 | -8.6566
4.8468 | -34.0983 | -8.6672
4.8529 | -34.0986 | -8.6779
4.859 | -34.0988 | -8.6885
4.8651 | -34.0991 | -8.6986
4.8712 | -34.0993 | -8.7088
4.8773 | -34.0995 | -8.7193
4.8834 | -34.0997 | -8.7296
4.8895 | -34.1 | -8.7397
4.8956 | -34.1002 | -8.7501
4.9017 | -34.1004 | -8.7603
4.9078 | -34.1006 | -8.771
4.9139 | -34.1009 | -8.781
4.92 | -34.1012 | -8.791
4.9261 | -34.1014 | -8.801
4.9322 | -34.1017 | -8.8113
4.9383 | -34.102 | -8.8216
4.9445 | -34.1022 | -8.8319
4.9506 | -34.1024 | -8.8421
4.9567 | -34.1027 | -8.8523
4.9628 | -34.1029 | -8.8626
4.9689 | -34.1032 | -8.8733
4.975 | -34.1034 | -8.8839
4.9811 | -34.1037 | -8.8939
4.9872 | -34.104 | -8.9041
4.9933 | -34.1043 | -8.9147
4.9994 | -34.1046 | -8.9254
5.0055 | -34.1049 | -8.9358
5.0116 | -34.1051 | -8.9465
5.0177 | -34.1055 | -8.9571
5.0238 | -34.1058 | -8.9678
5.0299 | -34.1061 | -8.9784
5.036 | -34.1064 | -8.9888
5.0421 | -34.1066 | -8.9991
5.0482 | -34.1069 | -9.0095
5.0543 | -34.1072 | -9.0197
5.0604 | -34.1075 | -9.03
5.0665 | -34.1078 | -9.0401
5.0726 | -34.1081 | -9.0508
5.0787 | -34.1084 | -9.0614
5.0848 | -34.1087 | -9.0719
5.091 | -34.1089 | -9.0821
5.0971 | -34.1093 | -9.0921
5.1032 | -34.1095 | -9.1029
5.1093 | -34.1098 | -9.1132
5.1154 | -34.1101 | -9.1234
5.1215 | -34.1103 | -9.1336
5.1276 | -34.1106 | -9.1438
5.1337 | -34.111 | -9.1539
5.1398 | -34.1112 | -9.1639
5.1459 | -34.1115 | -9.1746
5.152 | -34.1118 | -9.1851
5.1581 | -34.1121 | -9.1957
5.1642 | -34.1124 | -9.2061
5.1703 | -34.1127 | -9.2164
5.1764 | -34.113 | -9.2267
5.1825 | -34.1133 | -9.2368
5.1886 | -34.1136 | -9.2469
5.1947 | -34.1139 | -9.2572
5.2008 | -34.1142 | -9.2677
5.2069 | -34.1145 | -9.2782
5.213 | -34.1148 | -9.2883
5.2191 | -34.1151 | -9.2988
5.2252 | -34.1154 | -9.3092
5.2314 | -34.1157 | -9.3198
5.2375 | -34.116 | -9.3298
5.2436 | -34.1162 | -9.3402
5.2497 | -34.1165 | -9.3503
5.2558 | -34.1169 | -9.361
5.2619 | -34.1171 | -9.3717
5.268 | -34.1175 | -9.3822
5.2741 | -34.1178 | -9.3926
5.2802 | -34.1181 | -9.4034
5.2863 | -34.1185 | -9.4137
5.2924 | -34.1188 | -9.4242
5.2985 | -34.1191 | -9.4347
5.3046 | -34.1194 | -9.4451
5.3107 | -34.1197 | -9.4556
5.3168 | -34.12 | -9.466
5.3229 | -34.1203 | -9.4764
5.329 | -34.1206 | -9.4869
5.3351 | -34.1209 | -9.497
5.3412 | -34.1212 | -9.5073
5.3473 | -34.1215 | -9.5177
5.3534 | -34.1218 | -9.5282
5.3595 | -34.1221 | -9.5386
5.3656 | -34.1224 | -9.5489
5.3717 | -34.1227 | -9.5596
5.3779 | -34.123 | -9.5704
5.384 | -34.1233 | -9.5807
5.3901 | -34.1237 | -9.5914
5.3962 | -34.1239 | -9.6016
5.4023 | -34.1243 | -9.6122
5.4084 | -34.1246 | -9.6225
5.4145 | -34.1249 | -9.6328
5.4206 | -34.1252 | -9.6432
5.4267 | -34.1256 | -9.654
5.4328 | -34.1259 | -9.6644
5.4389 | -34.1262 | -9.6746
5.445 | -34.1264 | -9.6852
5.4511 | -34.1268 | -9.6955
5.4572 | -34.1271 | -9.706
5.4633 | -34.1274 | -9.7164
5.4694 | -34.1277 | -9.7268
5.4755 | -34.128 | -9.737
5.4816 | -34.1283 | -9.7475
5.4877 | -34.1286 | -9.7577
5.4938 | -34.1289 | -9.7679
5.4999 | -34.1292 | -9.7782
5.506 | -34.1296 | -9.7885
5.5121 | -34.1298 | -9.7988
5.5183 | -34.1301 | -9.8089
5.5244 | -34.1304 | -9.8195
5.5305 | -34.1308 | -9.8298
5.5366 | -34.1311 | -9.8399
5.5427 | -34.1313 | -9.8504
5.5488 | -34.1316 | -9.8608
5.5549 | -34.1319 | -9.8712
5.561 | -34.1322 | -9.8818
5.5671 | -34.1325 | -9.8921
5.5732 | -34.1328 | -9.9027
5.5793 | -34.1331 | -9.9133
5.5854 | -34.1335 | -9.9236
5.5915 | -34.1338 | -9.9343
5.5976 | -34.134 | -9.9447
5.6037 | -34.1343 | -9.9552
5.6098 | -34.1346 | -9.9655
5.6159 | -34.1349 | -9.9757
5.622 | -34.1352 | -9.9861
5.6281 | -34.1356 | -9.9965
5.6342 | -34.1359 | -10.0072
5.6403 | -34.1361 | -10.0176
5.6464 | -34.1364 | -10.0282
5.6525 | -34.1367 | -10.0387
5.6586 | -34.137 | -10.0488
5.6648 | -34.1374 | -10.0592
5.6709 | -34.1377 | -10.0698
5.677 | -34.138 | -10.0802
5.6831 | -34.1383 | -10.0907
5.6892 | -34.1385 | -10.1012
5.6953 | -34.1388 | -10.1115
5.7014 | -34.1392 | -10.122
5.7075 | -34.1394 | -10.1324
5.7136 | -34.1397 | -10.1427
5.7197 | -34.14 | -10.1532
5.7258 | -34.1402 | -10.1634
5.7319 | -34.1406 | -10.1738
5.738 | -34.1408 | -10.1842
5.7441 | -34.1412 | -10.1945
5.7502 | -34.1415 | -10.2047
5.7563 | -34.1418 | -10.215
5.7624 | -34.1421 | -10.2256
5.7685 | -34.1424 | -10.2361
5.7746 | -34.1427 | -10.2468
5.7807 | -34.1429 | -10.2572
5.7868 | -34.1432 | -10.2678
5.7929 | -34.1435 | -10.2781
5.799 | -34.1437 | -10.2883
5.8052 | -34.144 | -10.2985
5.8113 | -34.1443 | -10.3088
5.8174 | -34.1446 | -10.319
5.8235 | -34.1449 | -10.3294
5.8296 | -34.1452 | -10.34
5.8357 | -34.1454 | -10.3503
5.8418 | -34.1457 | -10.3608
5.8479 | -34.146 | -10.3712
5.854 | -34.1463 | -10.3818
5.8601 | -34.1466 | -10.3921
5.8662 | -34.1469 | -10.4026
5.8723 | -34.1471 | -10.4129
5.8784 | -34.1474 | -10.4231
5.8845 | -34.1477 | -10.4334
5.8906 | -34.148 | -10.4439
5.8967 | -34.1482 | -10.4541
5.9028 | -34.1486 | -10.4643
5.9089 | -34.1489 | -10.4745
5.915 | -34.1492 | -10.4847
5.9211 | -34.1494 | -10.495
5.9272 | -34.1497 | -10.5051
5.9333 | -34.15 | -10.5153
5.9394 | -34.1503 | -10.5256
5.9455 | -34.1506 | -10.5358
5.9517 | -34.151 | -10.546
5.9578 | -34.1513 | -10.5565
5.9639 | -34.1516 | -10.5665
5.97 | -34.1519 | -10.577
5.9761 | -34.1522 | -10.5872
5.9822 | -34.1525 | -10.5977
5.9883 | -34.1528 | -10.6081
5.9944 | -34.1531 | -10.6186
6.0005 | -34.1534 | -10.629
6.0066 | -34.1537 | -10.6392
6.0127 | -34.154 | -10.6493
6.0188 | -34.1543 | -10.6598
6.0249 | -34.1546 | -10.6703
6.031 | -34.1549 | -10.6804
6.0371 | -34.1552 | -10.6907
6.0432 | -34.1554 | -10.7009
6.0493 | -34.1558 | -10.7112
6.0554 | -34.1561 | -10.7213
6.0615 | -34.1564 | -10.7317
6.0676 | -34.1567 | -10.7422
6.0737 | -34.157 | -10.7526
6.0798 | -34.1573 | -10.7629
6.0859 | -34.1576 | -10.7732
6.0921 | -34.1579 | -10.7833
6.0982 | -34.1582 | -10.7936
6.1043 | -34.1585 | -10.8037
6.1104 | -34.1588 | -10.8143
6.1165 | -34.1591 | -10.8243
6.1226 | -34.1594 | -10.8345
6.1287 | -34.1597 | -10.8448
6.1348 | -34.1601 | -10.8551
6.1409 | -34.1604 | -10.8654
6.147 | -34.1607 | -10.8757
6.1531 | -34.161 | -10.8859
6.1592 | -34.1613 | -10.8961
6.1653 | -34.1615 | -10.9063
6.1714 | -34.1618 | -10.9167
6.1775 | -34.1622 | -10.9271
6.1836 | -34.1624 | -10.9374
6.1897 | -34.1628 | -10.9477
6.1958 | -34.163 | -10.9582
6.2019 | -34.1634 | -10.9686
6.208 | -34.1637 | -10.9789
6.2141 | -34.164 | -10.9891
6.2202 | -34.1643 | -10.9995
6.2263 | -34.1646 | -11.0098
6.2325 | -34.1649 | -11.0198
6.2386 | -34.1652 | -11.0299
6.2447 | -34.1655 | -11.0404
6.2508 | -34.1659 | -11.0507
6.2569 | -34.1662 | -11.061
6.263 | -34.1665 | -11.0712
6.2691 | -34.1668 | -11.0815
6.2752 | -34.167 | -11.0919
6.2813 | -34.1673 | -11.1022
6.2874 | -34.1676 | -11.1123
6.2935 | -34.168 | -11.1226
6.2996 | -34.1682 | -11.1328
6.3057 | -34.1685 | -11.1432
6.3118 | -34.1688 | -11.1536
6.3179 | -34.1691 | -11.1638
6.324 | -34.1694 | -11.1742
6.3301 | -34.1697 | -11.1845
6.3362 | -34.17 | -11.1946
6.3423 | -34.1704 | -11.2051
6.3484 | -34.1707 | -11.2155
6.3545 | -34.171 | -11.2255
6.3606 | -34.1713 | -11.2355
6.3667 | -34.1715 | -11.2457
6.3728 | -34.1719 | -11.256
6.379 | -34.1722 | -11.2663
6.3851 | -34.1725 | -11.2763
6.3912 | -34.1727 | -11.2865
6.3973 | -34.173 | -11.2969
6.4034 | -34.1733 | -11.307
6.4095 | -34.1736 | -11.3173
6.4156 | -34.1739 | -11.3276
6.4217 | -34.1742 | -11.3378
6.4278 | -34.1745 | -11.3482
6.4339 | -34.1749 | -11.3584
6.44 | -34.1752 | -11.3686
6.4461 | -34.1755 | -11.379
6.4522 | -34.1758 | -11.3891
6.4583 | -34.1761 | -11.3994
6.4644 | -34.1764 | -11.4098
6.4705 | -34.1767 | -11.4198
6.4766 | -34.177 | -11.43
6.4827 | -34.1773 | -11.4403
6.4888 | -34.1776 | -11.4504
6.4949 | -34.1779 | -11.4604
6.501 | -34.1782 | -11.4707
6.5071 | -34.1785 | -11.4808
6.5132 | -34.1788 | -11.4911
6.5194 | -34.1791 | -11.5014
6.5255 | -34.1794 | -11.5114
6.5316 | -34.1797 | -11.5217
6.5377 | -34.18 | -11.532
6.5438 | -34.1803 | -11.5421
6.5499 | -34.1806 | -11.5523
6.556 | -34.1809 | -11.5625
6.5621 | -34.1812 | -11.5729
6.5682 | -34.1815 | -11.5833
6.5743 | -34.1819 | -11.5934
6.5804 | -34.1822 | -11.6038
6.5865 | -34.1825 | -11.6141
6.5926 | -34.1828 | -11.6243
6.5987 | -34.1831 | -11.6342
6.6048 | -34.1834 | -11.6445
6.6109 | -34.1837 | -11.6548
6.617 | -34.184 | -11.6652
6.6231 | -34.1843 | -11.6754
6.6292 | -34.1846 | -11.6857
6.6353 | -34.1849 | -11.6961
6.6414 | -34.1853 | -11.7062
6.6475 | -34.1856 | -11.7165
6.6536 | -34.1859 | -11.7267
6.6597 | -34.1862 | -11.7367
6.6659 | -34.1865 | -11.7469
6.672 | -34.1868 | -11.7571
6.6781 | -34.1871 | -11.7674
6.6842 | -34.1874 | -11.7775
6.6903 | -34.1878 | -11.7875
6.6964 | -34.1881 | -11.7977
6.7025 | -34.1884 | -11.8078
6.7086 | -34.1887 | -11.818
6.7147 | -34.189 | -11.8283
6.7208 | -34.1893 | -11.8386
6.7269 | -34.1896 | -11.8487
6.733 | -34.1899 | -11.8587
6.7391 | -34.1903 | -11.869
6.7452 | -34.1906 | -11.879
6.7513 | -34.1909 | -11.889
6.7574 | -34.1912 | -11.8993
6.7635 | -34.1915 | -11.9093
6.7696 | -34.1918 | -11.9194
6.7757 | -34.1922 | -11.9295
6.7818 | -34.1925 | -11.9398
6.7879 | -34.1929 | -11.9498
6.794 | -34.1932 | -11.9598
6.8001 | -34.1934 | -11.97
6.8063 | -34.1938 | -11.9801
6.8124 | -34.1941 | -11.9904
6.8185 | -34.1944 | -12.0008
6.8246 | -34.1947 | -12.0108
6.8307 | -34.195 | -12.0209
6.8368 | -34.1953 | -12.0313
6.8429 | -34.1956 | -12.0415
6.849 | -34.196 | -12.0517
6.8551 | -34.1963 | -12.0617
6.8612 | -34.1966 | -12.0719
6.8673 | -34.1969 | -12.082
6.8734 | -34.1972 | -12.0923
6.8795 | -34.1976 | -12.1024
6.8856 | -34.1979 | -12.1126
6.8917 | -34.1982 | -12.1228
6.8978 | -34.1985 | -12.1327
6.9039 | -34.1988 | -12.1426
6.91 | -34.1992 | -12.1528
6.9161 | -34.1995 | -12.1627
6.9222 | -34.1998 | -12.1729
6.9283 | -34.2002 | -12.1827
6.9344 | -34.2005 | -12.1929
6.9405 | -34.2009 | -12.2029
6.9466 | -34.2012 | -12.2132
6.9528 | -34.2015 | -12.2235
6.9589 | -34.2019 | -12.2337
6.965 | -34.2022 | -12.244
6.9711 | -34.2026 | -12.2542
6.9772 | -34.2029 | -12.2643
6.9833 | -34.2032 | -12.2744
6.9894 | -34.2036 | -12.2847
6.9955 | -34.2039 | -12.2946
7.0016 | -34.2042 | -12.3047
7.0077 | -34.2046 | -12.3148
7.0138 | -34.2049 | -12.3249
7.0199 | -34.2052 | -12.3349
7.026 | -34.2056 | -12.345
7.0321 | -34.2059 | -12.3548
7.0382 | -34.2062 | -12.3647
7.0443 | -34.2066 | -12.3749
7.0504 | -34.2069 | -12.385
7.0565 | -34.2072 | -12.3951
7.0626 | -34.2076 | -12.405
7.0687 | -34.2079 | -12.4147
7.0748 | -34.2083 | -12.4245
7.0809 | -34.2086 | -12.4345
7.087 | -34.2089 | -12.4447
7.0932 | -34.2093 | -12.4547
7.0993 | -34.2096 | -12.4648
7.1054 | -34.2099 | -12.4747
7.1115 | -34.2103 | -12.4846
7.1176 | -34.2106 | -12.4945
7.1237 | -34.2109 | -12.5045
7.1298 | -34.2113 | -12.5144
7.1359 | -34.2116 | -12.5244
7.142 | -34.2119 | -12.5342
7.1481 | -34.2122 | -12.5442
7.1542 | -34.2126 | -12.5542
7.1603 | -34.2129 | -12.5642
7.1664 | -34.2133 | -12.5742
7.1725 | -34.2136 | -12.5841
7.1786 | -34.214 | -12.5941
7.1847 | -34.2143 | -12.6042
7.1908 | -34.2146 | -12.6143
7.1969 | -34.2149 | -12.6245
7.203 | -34.2153 | -12.6343
7.2091 | -34.2157 | -12.6441
7.2152 | -34.216 | -12.6541
7.2213 | -34.2163 | -12.6639
7.2274 | -34.2167 | -12.6739
7.2335 | -34.217 | -12.6838
7.2397 | -34.2174 | -12.6937
7.2458 | -34.2177 | -12.7034
7.2519 | -34.2181 | -12.7134
7.258 | -34.2184 | -12.7236
7.2641 | -34.2188 | -12.7335
7.2702 | -34.2191 | -12.7437
7.2763 | -34.2194 | -12.7536
7.2824 | -34.2198 | -12.7636
7.2885 | -34.2201 | -12.7736
7.2946 | -34.2205 | -12.7835
7.3007 | -34.2208 | -12.7936
7.3068 | -34.2211 | -12.8038
7.3129 | -34.2215 | -12.8137
7.319 | -34.2219 | -12.8237
7.3251 | -34.2222 | -12.8335
7.3312 | -34.2225 | -12.8435
7.3373 | -34.2229 | -12.8532
7.3434 | -34.2233 | -12.863
7.3495 | -34.2236 | -12.8728
7.3556 | -34.224 | -12.8828
7.3617 | -34.2243 | -12.8929
7.3678 | -34.2247 | -12.9028
7.3739 | -34.225 | -12.9129
7.3801 | -34.2253 | -12.9231
7.3862 | -34.2257 | -12.9329
7.3923 | -34.226 | -12.943
7.3984 | -34.2264 | -12.9527
7.4045 | -34.2267 | -12.9625
7.4106 | -34.2271 | -12.9724
7.4167 | -34.2274 | -12.9824
7.4228 | -34.2278 | -12.9924
7.4289 | -34.2281 | -13.0023
7.435 | -34.2285 | -13.012
7.4411 | -34.2289 | -13.0219
7.4472 | -34.2292 | -13.0319
7.4533 | -34.2295 | -13.0416
7.4594 | -34.2299 | -13.0517
7.4655 | -34.2303 | -13.0614
7.4716 | -34.2306 | -13.0715
7.4777 | -34.231 | -13.0814
7.4838 | -34.2313 | -13.0912
7.4899 | -34.2316 | -13.1009
7.496 | -34.232 | -13.1106
7.5021 | -34.2324 | -13.1205
7.5082 | -34.2327 | -13.1305
7.5143 | -34.233 | -13.1404
7.5204 | -34.2334 | -13.1504
7.5266 | -34.2337 | -13.1602
7.5327 | -34.2341 | -13.1701
7.5388 | -34.2344 | -13.1802
7.5449 | -34.2348 | -13.1902
7.551 | -34.2351 | -13.2
7.5571 | -34.2355 | -13.21
7.5632 | -34.2359 | -13.2199
7.5693 | -34.2362 | -13.2297
7.5754 | -34.2366 | -13.2395
7.5815 | -34.2369 | -13.2494
7.5876 | -34.2373 | -13.2595
7.5937 | -34.2376 | -13.2694
7.5998 | -34.238 | -13.2793
7.6059 | -34.2384 | -13.2894
7.612 | -34.2387 | -13.2993
7.6181 | -34.2391 | -13.309
7.6242 | -34.2394 | -13.3192
7.6303 | -34.2398 | -13.3291
7.6364 | -34.2401 | -13.3391
7.6425 | -34.2405 | -13.349
7.6486 | -34.2408 | -13.3589
7.6547 | -34.2412 | -13.3687
7.6608 | -34.2415 | -13.3786
7.667 | -34.2419 | -13.3885
7.6731 | -34.2422 | -13.3985
7.6792 | -34.2426 | -13.4086
7.6853 | -34.2429 | -13.4186
7.6914 | -34.2433 | -13.4286
7.6975 | -34.2436 | -13.4386
7.7036 | -34.244 | -13.4485
7.7097 | -34.2443 | -13.4586
7.7158 | -34.2447 | -13.4685
7.7219 | -34.2451 | -13.4783
7.728 | -34.2454 | -13.4882
7.7341 | -34.2458 | -13.4982
7.7402 | -34.2461 | -13.5082
7.7463 | -34.2465 | -13.5181
7.7524 | -34.2468 | -13.5283
7.7585 | -34.2472 | -13.5382
7.7646 | -34.2475 | -13.5481
7.7707 | -34.2478 | -13.5581
7.7768 | -34.2482 | -13.5682
7.7829 | -34.2485 | -13.5784
7.789 | -34.2489 | -13.5886
7.7951 | -34.2492 | -13.5985
7.8012 | -34.2496 | -13.6085
7.8073 | -34.25 | -13.6186
7.8135 | -34.2503 | -13.6285
7.8196 | -34.2507 | -13.6386
7.8257 | -34.2511 | -13.6488
7.8318 | -34.2514 | -13.6588
7.8379 | -34.2518 | -13.6687
7.844 | -34.2521 | -13.6786
7.8501 | -34.2525 | -13.6887
7.8562 | -34.2528 | -13.6989
7.8623 | -34.2532 | -13.7088
7.8684 | -34.2535 | -13.7188
7.8745 | -34.2538 | -13.7289
7.8806 | -34.2542 | -13.739
7.8867 | -34.2546 | -13.7488
7.8928 | -34.2549 | -13.7585
7.8989 | -34.2553 | -13.7685
7.905 | -34.2556 | -13.7785
7.9111 | -34.256 | -13.7885
7.9172 | -34.2563 | -13.7984
7.9233 | -34.2567 | -13.8084
7.9294 | -34.257 | -13.8185
7.9355 | -34.2574 | -13.8283
7.9416 | -34.2577 | -13.8383
7.9477 | -34.258 | -13.8482
7.9539 | -34.2584 | -13.8581
7.96 | -34.2587 | -13.8681
7.9661 | -34.2591 | -13.8781
7.9722 | -34.2595 | -13.8879
7.9783 | -34.2598 | -13.8981
7.9844 | -34.2602 | -13.9079
7.9905 | -34.2605 | -13.9181
7.9966 | -34.2609 | -13.9281
8.0027 | -34.2612 | -13.9381
8.0088 | -34.2615 | -13.9481
8.0149 | -34.2619 | -13.958
8.021 | -34.2622 | -13.968
8.0271 | -34.2626 | -13.9782
8.0332 | -34.2629 | -13.9881
8.0393 | -34.2633 | -13.9981
8.0454 | -34.2637 | -14.0083
8.0515 | -34.264 | -14.0183
8.0576 | -34.2644 | -14.0285
8.0637 | -34.2647 | -14.0384
8.0698 | -34.2651 | -14.0482
8.0759 | -34.2654 | -14.0581
8.082 | -34.2657 | -14.068
8.0881 | -34.2661 | -14.0781
8.0942 | -34.2664 | -14.0881
8.1004 | -34.2668 | -14.0982
8.1065 | -34.2672 | -14.1081
8.1126 | -34.2675 | -14.1181
8.1187 | -34.2678 | -14.1282
8.1248 | -34.2682 | -14.1379
8.1309 | -34.2685 | -14.1478
8.137 | -34.2689 | -14.1575
8.1431 | -34.2692 | -14.1677
8.1492 | -34.2696 | -14.1777
8.1553 | -34.2699 | -14.1877
8.1614 | -34.2703 | -14.1979
8.1675 | -34.2706 | -14.2079
8.1736 | -34.2709 | -14.218
8.1797 | -34.2713 | -14.2281
8.1858 | -34.2716 | -14.2381
8.1919 | -34.272 | -14.248
8.198 | -34.2723 | -14.2579
8.2041 | -34.2727 | -14.2679
8.2102 | -34.273 | -14.2778
8.2163 | -34.2734 | -14.2876
8.2224 | -34.2737 | -14.2976
8.2285 | -34.2741 | -14.3074
8.2346 | -34.2744 | -14.3173
8.2408 | -34.2748 | -14.3272
8.2469 | -34.2751 | -14.3372
8.253 | -34.2755 | -14.347
8.2591 | -34.2759 | -14.3571
8.2652 | -34.2762 | -14.3671
8.2713 | -34.2766 | -14.3772
8.2774 | -34.2769 | -14.3872
8.2835 | -34.2773 | -14.3974
8.2896 | -34.2776 | -14.4072
8.2957 | -34.278 | -14.4172
8.3018 | -34.2783 | -14.4272
8.3079 | -34.2787 | -14.4371
8.314 | -34.2791 | -14.4471
8.3201 | -34.2794 | -14.457
8.3262 | -34.2797 | -14.467
8.3323 | -34.2801 | -14.4769
8.3384 | -34.2805 | -14.4869
8.3445 | -34.2808 | -14.4968
8.3506 | -34.2812 | -14.5068
8.3567 | -34.2815 | -14.5169
8.3628 | -34.2819 | -14.5269
8.3689 | -34.2822 | -14.5368
8.375 | -34.2825 | -14.5468
8.3812 | -34.2829 | -14.5566
8.3873 | -34.2832 | -14.5665
8.3934 | -34.2836 | -14.5765
8.3995 | -34.2839 | -14.5863
8.4056 | -34.2843 | -14.5963
8.4117 | -34.2846 | -14.6063
8.4178 | -34.285 | -14.6163
8.4239 | -34.2853 | -14.6261
8.43 | -34.2856 | -14.6361
8.4361 | -34.2859 | -14.6461
8.4422 | -34.2863 | -14.656
8.4483 | -34.2867 | -14.6658
8.4544 | -34.287 | -14.6758
8.4605 | -34.2874 | -14.6856
8.4666 | -34.2877 | -14.6957
8.4727 | -34.2881 | -14.7057
8.4788 | -34.2884 | -14.7156
8.4849 | -34.2888 | -14.7255
8.491 | -34.2891 | -14.7354
8.4971 | -34.2895 | -14.7453
8.5032 | -34.2898 | -14.7551
8.5093 | -34.2902 | -14.7649
8.5154 | -34.2905 | -14.7748
8.5215 | -34.2909 | -14.7846
8.5277 | -34.2912 | -14.7946
8.5338 | -34.2916 | -14.8045
8.5399 | -34.2919 | -14.8144
8.546 | -34.2923 | -14.8241
8.5521 | -34.2927 | -14.834
8.5582 | -34.293 | -14.8439
8.5643 | -34.2933 | -14.8536
8.5704 | -34.2937 | -14.8633
8.5765 | -34.294 | -14.8733
8.5826 | -34.2944 | -14.8829
8.5887 | -34.2948 | -14.8927
8.5948 | -34.2951 | -14.9026
8.6009 | -34.2955 | -14.9123
8.607 | -34.2958 | -14.9222
8.6131 | -34.2962 | -14.9321
8.6192 | -34.2966 | -14.9419
8.6253 | -34.2969 | -14.9517
8.6314 | -34.2973 | -14.9615
8.6375 | -34.2976 | -14.9714
8.6436 | -34.298 | -14.9812
8.6497 | -34.2983 | -14.991
8.6558 | -34.2987 | -15.0009
8.6619 | -34.299 | -15.0108
8.6681 | -34.2994 | -15.0208
8.6742 | -34.2998 | -15.0307
8.6803 | -34.3001 | -15.0404
8.6864 | -34.3005 | -15.05
8.6925 | -34.3008 | -15.0597
8.6986 | -34.3012 | -15.0696
8.7047 | -34.3015 | -15.0796
8.7108 | -34.3019 | -15.0894
8.7169 | -34.3022 | -15.0993
8.723 | -34.3026 | -15.1089
8.7291 | -34.3029 | -15.1186
8.7352 | -34.3033 | -15.1285
8.7413 | -34.3036 | -15.1385
8.7474 | -34.304 | -15.1481
8.7535 | -34.3044 | -15.158
8.7596 | -34.3047 | -15.1678
8.7657 | -34.3051 | -15.1778
8.7718 | -34.3054 | -15.1877
8.7779 | -34.3058 | -15.1975
8.784 | -34.3062 | -15.2074
8.7901 | -34.3066 | -15.2172
8.7962 | -34.3069 | -15.2268
8.8023 | -34.3073 | -15.2365
8.8084 | -34.3076 | -15.2462
8.8146 | -34.308 | -15.256
8.8207 | -34.3083 | -15.2658
8.8268 | -34.3087 | -15.2753
8.8329 | -34.3091 | -15.2852
8.839 | -34.3095 | -15.295
8.8451 | -34.3098 | -15.3047
8.8512 | -34.3102 | -15.3143
8.8573 | -34.3105 | -15.3241
8.8634 | -34.3109 | -15.3337
8.8695 | -34.3113 | -15.3437
8.8756 | -34.3116 | -15.3535
8.8817 | -34.312 | -15.3631
8.8878 | -34.3123 | -15.3731
8.8939 | -34.3127 | -15.3829
8.9 | -34.3131 | -15.3925
8.9061 | -34.3134 | -15.4023
8.9122 | -34.3138 | -15.4121
8.9183 | -34.3142 | -15.4219
8.9244 | -34.3146 | -15.4319
8.9305 | -34.315 | -15.4414
8.9366 | -34.3153 | -15.4511
8.9427 | -34.3157 | -15.4609
8.9488 | -34.3161 | -15.4706
8.955 | -34.3164 | -15.4803
8.9611 | -34.3168 | -15.4902
8.9672 | -34.3172 | -15.4999
8.9733 | -34.3176 | -15.5095
8.9794 | -34.318 | -15.5192
8.9855 | -34.3183 | -15.5291
8.9916 | -34.3187 | -15.5389
8.9977 | -34.3191 | -15.5487
9.0038 | -34.3195 | -15.5583
9.0099 | -34.3199 | -15.568
9.016 | -34.3202 | -15.5778
9.0221 | -34.3206 | -15.5873
9.0282 | -34.321 | -15.5969
9.0343 | -34.3214 | -15.6066
9.0404 | -34.3218 | -15.6163
9.0465 | -34.3221 | -15.6263
9.0526 | -34.3225 | -15.6359
9.0587 | -34.3229 | -15.6456
9.0648 | -34.3233 | -15.6552
9.0709 | -34.3237 | -15.6649
9.077 | -34.3241 | -15.6747
9.0831 | -34.3244 | -15.6845
9.0892 | -34.3248 | -15.6943
9.0953 | -34.3252 | -15.7041
9.1015 | -34.3256 | -15.7138
9.1076 | -34.326 | -15.7232
9.1137 | -34.3264 | -15.7329
9.1198 | -34.3268 | -15.7425
9.1259 | -34.3272 | -15.7521
9.132 | -34.3275 | -15.762
9.1381 | -34.3279 | -15.7717
9.1442 | -34.3284 | -15.7811
9.1503 | -34.3287 | -15.7909
9.1564 | -34.3291 | -15.8005
9.1625 | -34.3295 | -15.8104
9.1686 | -34.3299 | -15.82
9.1747 | -34.3303 | -15.8297
9.1808 | -34.3307 | -15.8393
9.1869 | -34.3311 | -15.849
9.193 | -34.3314 | -15.8587
9.1991 | -34.3318 | -15.8685
9.2052 | -34.3322 | -15.8782
9.2113 | -34.3326 | -15.8879
9.2174 | -34.333 | -15.8976
9.2235 | -34.3334 | -15.9073
9.2296 | -34.3338 | -15.9169
9.2357 | -34.3342 | -15.9267
9.2419 | -34.3346 | -15.9363
9.248 | -34.3351 | -15.946
9.2541 | -34.3354 | -15.9558
9.2602 | -34.3358 | -15.9655
9.2663 | -34.3362 | -15.975
9.2724 | -34.3366 | -15.9846
9.2785 | -34.337 | -15.9943
9.2846 | -34.3374 | -16.0039
9.2907 | -34.3378 | -16.0135
9.2968 | -34.3382 | -16.0233
9.3029 | -34.3386 | -16.033
9.309 | -34.339 | -16.0427
9.3151 | -34.3394 | -16.0522
9.3212 | -34.3398 | -16.0617
9.3273 | -34.3402 | -16.0714
9.3334 | -34.3406 | -16.0809
9.3395 | -34.341 | -16.0903
9.3456 | -34.3414 | -16.0999
9.3517 | -34.3418 | -16.1096
9.3578 | -34.3422 | -16.1192
9.3639 | -34.3426 | -16.1287
9.37 | -34.343 | -16.1383
9.3761 | -34.3434 | -16.1481
9.3822 | -34.3438 | -16.1578
9.3884 | -34.3442 | -16.1672
9.3945 | -34.3446 | -16.1769
9.4006 | -34.3451 | -16.1865
9.4067 | -34.3454 | -16.1961
9.4128 | -34.3458 | -16.2058
9.4189 | -34.3463 | -16.2156
9.425 | -34.3467 | -16.2251
9.4311 | -34.3471 | -16.2347
9.4372 | -34.3475 | -16.2444
9.4433 | -34.3479 | -16.254
9.4494 | -34.3483 | -16.2635
9.4555 | -34.3487 | -16.2732
9.4616 | -34.3491 | -16.2828
9.4677 | -34.3496 | -16.2925
9.4738 | -34.35 | -16.3021
9.4799 | -34.3504 | -16.3118
9.486 | -34.3508 | -16.3214
9.4921 | -34.3512 | -16.331
9.4982 | -34.3516 | -16.3403
9.5043 | -34.352 | -16.35
9.5104 | -34.3524 | -16.3594
9.5165 | -34.3528 | -16.3691
9.5226 | -34.3532 | -16.3788
9.5288 | -34.3536 | -16.3883
9.5349 | -34.354 | -16.398
9.541 | -34.3544 | -16.4074
9.5471 | -34.3548 | -16.417
9.5532 | -34.3553 | -16.4267
9.5593 | -34.3557 | -16.4364
9.5654 | -34.3561 | -16.4459
9.5715 | -34.3565 | -16.4555
9.5776 | -34.3569 | -16.4652
9.5837 | -34.3573 | -16.4747
9.5898 | -34.3577 | -16.4844
9.5959 | -34.3581 | -16.4941
9.602 | -34.3585 | -16.5035
9.6081 | -34.3589 | -16.5131
9.6142 | -34.3593 | -16.5226
9.6203 | -34.3597 | -16.5323
9.6264 | -34.3601 | -16.5418
9.6325 | -34.3605 | -16.5515
9.6386 | -34.3609 | -16.5611
9.6447 | -34.3613 | -16.5707
9.6508 | -34.3617 | -16.5802
9.6569 | -34.3621 | -16.5898
9.663 | -34.3626 | -16.5995
9.6691 | -34.363 | -16.609
9.6753 | -34.3634 | -16.6185
9.6814 | -34.3638 | -16.6279
9.6875 | -34.3643 | -16.6373
9.6936 | -34.3647 | -16.6467
9.6997 | -34.3651 | -16.6564
9.7058 | -34.3655 | -16.666
9.7119 | -34.3659 | -16.6756
9.718 | -34.3663 | -16.6853
9.7241 | -34.3668 | -16.6948
9.7302 | -34.3671 | -16.7044
9.7363 | -34.3676 | -16.7138
9.7424 | -34.368 | -16.7233
9.7485 | -34.3684 | -16.7329
9.7546 | -34.3688 | -16.7423
9.7607 | -34.3692 | -16.7519
9.7668 | -34.3696 | -16.7616
9.7729 | -34.3701 | -16.7711
9.779 | -34.3705 | -16.7805
9.7851 | -34.3709 | -16.7902
9.7912 | -34.3714 | -16.7997
9.7973 | -34.3718 | -16.8093
9.8034 | -34.3722 | -16.8187
9.8095 | -34.3726 | -16.8285
9.8157 | -34.3731 | -16.838
9.8218 | -34.3735 | -16.8477
9.8279 | -34.3739 | -16.8571
9.834 | -34.3742 | -16.8667
9.8401 | -34.3747 | -16.8763
9.8462 | -34.3751 | -16.8856
9.8523 | -34.3755 | -16.8951
9.8584 | -34.376 | -16.9046
9.8645 | -34.3764 | -16.9143
9.8706 | -34.3768 | -16.924
9.8767 | -34.3772 | -16.9334
9.8828 | -34.3776 | -16.9428
9.8889 | -34.378 | -16.9522
9.895 | -34.3784 | -16.9618
9.9011 | -34.3788 | -16.9713
9.9072 | -34.3792 | -16.9808
9.9133 | -34.3796 | -16.9903
9.9194 | -34.3801 | -16.9998
9.9255 | -34.3805 | -17.0092
9.9316 | -34.3809 | -17.0188
9.9377 | -34.3813 | -17.0282
9.9438 | -34.3817 | -17.0378
9.9499 | -34.3821 | -17.0476
9.956 | -34.3825 | -17.0571
9.9622 | -34.383 | -17.0667
9.9683 | -34.3834 | -17.0763
9.9744 | -34.3838 | -17.086
9.9805 | -34.3842 | -17.0955
9.9866 | -34.3846 | -17.1049
9.9927 | -34.385 | -17.1144
9.9988 | -34.3854 | -17.1241
10.0049 | -34.3859 | -17.1335
10.011 | -34.3863 | -17.1431
10.0171 | -34.3867 | -17.1526
10.0232 | -34.3871 | -17.162
10.0293 | -34.3876 | -17.1715
10.0354 | -34.388 | -17.1811
10.0415 | -34.3884 | -17.1905
10.0476 | -34.3888 | -17.2001
10.0537 | -34.3892 | -17.2097
10.0598 | -34.3896 | -17.2192
10.0659 | -34.39 | -17.2288
10.072 | -34.3905 | -17.2381
10.0781 | -34.3909 | -17.2476
10.0842 | -34.3913 | -17.2572
10.0903 | -34.3917 | -17.2667
10.0964 | -34.3921 | -17.276
10.1026 | -34.3925 | -17.2855
10.1087 | -34.3929 | -17.2949
10.1148 | -34.3933 | -17.3045
10.1209 | -34.3937 | -17.3139
10.127 | -34.3942 | -17.3232
10.1331 | -34.3946 | -17.3327
10.1392 | -34.395 | -17.3421
10.1453 | -34.3954 | -17.3516
10.1514 | -34.3958 | -17.3611
10.1575 | -34.3962 | -17.3707
10.1636 | -34.3967 | -17.3799
10.1697 | -34.3971 | -17.3894
10.1758 | -34.3975 | -17.3988
10.1819 | -34.3979 | -17.4082
10.188 | -34.3983 | -17.4176
10.1941 | -34.3987 | -17.4271
10.2002 | -34.3991 | -17.4366
10.2063 | -34.3995 | -17.446
10.2124 | -34.4 | -17.4555
10.2185 | -34.4004 | -17.4651
10.2246 | -34.4008 | -17.4745
10.2307 | -34.4012 | -17.484
10.2368 | -34.4016 | -17.4932
10.2429 | -34.402 | -17.5027
10.2491 | -34.4025 | -17.5123
10.2552 | -34.4029 | -17.5218
10.2613 | -34.4033 | -17.5313
10.2674 | -34.4037 | -17.5408
10.2735 | -34.4041 | -17.5505
10.2796 | -34.4046 | -17.5599
10.2857 | -34.405 | -17.5695
10.2918 | -34.4054 | -17.579
10.2979 | -34.4058 | -17.5884
10.304 | -34.4062 | -17.5976
10.3101 | -34.4066 | -17.6072
10.3162 | -34.4071 | -17.6168
10.3223 | -34.4075 | -17.6261
10.3284 | -34.4079 | -17.6355
10.3345 | -34.4083 | -17.6449
10.3406 | -34.4087 | -17.6542
10.3467 | -34.4091 | -17.6636
10.3528 | -34.4095 | -17.6729
10.3589 | -34.4099 | -17.6824
10.365 | -34.4104 | -17.692
10.3711 | -34.4108 | -17.7016
10.3772 | -34.4112 | -17.711
10.3833 | -34.4116 | -17.7204
10.3895 | -34.4121 | -17.7298
10.3956 | -34.4125 | -17.7394
10.4017 | -34.4129 | -17.7487
10.4078 | -34.4133 | -17.7582
10.4139 | -34.4137 | -17.7675
10.42 | -34.4142 | -17.7771
10.4261 | -34.4146 | -17.7864
10.4322 | -34.415 | -17.7958
10.4383 | -34.4154 | -17.8051
10.4444 | -34.4158 | -17.8145
10.4505 | -34.4162 | -17.8239
10.4566 | -34.4166 | -17.8334
10.4627 | -34.417 | -17.8427
10.4688 | -34.4174 | -17.852
10.4749 | -34.4178 | -17.8616
10.481 | -34.4183 | -17.8709
10.4871 | -34.4187 | -17.8801
10.4932 | -34.4191 | -17.8894
10.4993 | -34.4195 | -17.8988
10.5054 | -34.42 | -17.9082
10.5115 | -34.4204 | -17.9177
10.5176 | -34.4208 | -17.927
10.5237 | -34.4213 | -17.9366
10.5298 | -34.4216 | -17.9458
10.536 | -34.422 | -17.9552
10.5421 | -34.4225 | -17.9646
10.5482 | -34.4228 | -17.9742
10.5543 | -34.4233 | -17.9836
10.5604 | -34.4237 | -17.9929
10.5665 | -34.4241 | -18.0021
10.5726 | -34.4245 | -18.0116
10.5787 | -34.4249 | -18.021
10.5848 | -34.4253 | -18.0302
10.5909 | -34.4258 | -18.0398
10.597 | -34.4262 | -18.0493
10.6031 | -34.4266 | -18.0588
10.6092 | -34.427 | -18.0682
10.6153 | -34.4274 | -18.0776
10.6214 | -34.4278 | -18.0868
10.6275 | -34.4282 | -18.0961
10.6336 | -34.4286 | -18.1056
10.6397 | -34.429 | -18.1151
10.6458 | -34.4295 | -18.1246
10.6519 | -34.4299 | -18.134
10.658 | -34.4303 | -18.1434
10.6641 | -34.4307 | -18.1531
10.6702 | -34.4312 | -18.1625
10.6764 | -34.4316 | -18.172
10.6825 | -34.432 | -18.1811
10.6886 | -34.4324 | -18.1904
10.6947 | -34.4328 | -18.1996
10.7008 | -34.4333 | -18.209
10.7069 | -34.4336 | -18.2184
10.713 | -34.4341 | -18.2279
10.7191 | -34.4345 | -18.2373
10.7252 | -34.4349 | -18.2465
10.7313 | -34.4353 | -18.2557
10.7374 | -34.4357 | -18.265
10.7435 | -34.4361 | -18.2744
10.7496 | -34.4366 | -18.2834
10.7557 | -34.437 | -18.2929
10.7618 | -34.4374 | -18.3019
10.7679 | -34.4378 | -18.3114
10.774 | -34.4382 | -18.3207
10.7801 | -34.4386 | -18.3301
10.7862 | -34.439 | -18.3393
10.7923 | -34.4394 | -18.3488
10.7984 | -34.4399 | -18.3579
10.8045 | -34.4403 | -18.3673
10.8106 | -34.4407 | -18.3767
10.8168 | -34.4411 | -18.3859
10.8229 | -34.4415 | -18.3953
10.829 | -34.4419 | -18.4047
10.8351 | -34.4423 | -18.4137
10.8412 | -34.4428 | -18.4229
10.8473 | -34.4432 | -18.4322
10.8534 | -34.4436 | -18.4416
10.8595 | -34.444 | -18.451
10.8656 | -34.4444 | -18.4601
10.8717 | -34.4448 | -18.4695
10.8778 | -34.4453 | -18.4786
10.8839 | -34.4457 | -18.4879
10.89 | -34.4461 | -18.4973
10.8961 | -34.4465 | -18.5067
10.9022 | -34.4469 | -18.5159
10.9083 | -34.4473 | -18.5252
10.9144 | -34.4477 | -18.5344
10.9205 | -34.4481 | -18.5435
10.9266 | -34.4485 | -18.5527
10.9327 | -34.449 | -18.5619
10.9388 | -34.4494 | -18.571
10.9449 | -34.4498 | -18.5801
10.951 | -34.4502 | -18.5895
10.9571 | -34.4506 | -18.5987
10.9633 | -34.451 | -18.6079
10.9694 | -34.4515 | -18.6174
10.9755 | -34.4519 | -18.6268
10.9816 | -34.4523 | -18.6363
10.9877 | -34.4528 | -18.6455
10.9938 | -34.4532 | -18.6549
10.9999 | -34.4537 | -18.6641
11.006 | -34.4541 | -18.6731
11.0121 | -34.4545 | -18.6823
11.0182 | -34.4549 | -18.6915
11.0243 | -34.4553 | -18.7008
11.0304 | -34.4557 | -18.7102
11.0365 | -34.4561 | -18.7195
11.0426 | -34.4565 | -18.7287
11.0487 | -34.457 | -18.7381
11.0548 | -34.4574 | -18.7474
11.0609 | -34.4579 | -18.7567
11.067 | -34.4583 | -18.7661
11.0731 | -34.4587 | -18.7753
11.0792 | -34.4592 | -18.7845
11.0853 | -34.4596 | -18.7938
11.0914 | -34.46 | -18.8031
11.0975 | -34.4604 | -18.8125
11.1037 | -34.4609 | -18.8217
11.1098 | -34.4613 | -18.8313
11.1159 | -34.4617 | -18.8404
11.122 | -34.4621 | -18.8496
11.1281 | -34.4625 | -18.8589
11.1342 | -34.4629 | -18.8683
11.1403 | -34.4633 | -18.8775
11.1464 | -34.4638 | -18.8867
11.1525 | -34.4642 | -18.8955
11.1586 | -34.4646 | -18.9048
11.1647 | -34.465 | -18.9138
11.1708 | -34.4654 | -18.9232
11.1769 | -34.4658 | -18.9324
11.183 | -34.4663 | -18.9416
11.1891 | -34.4668 | -18.951
11.1952 | -34.4672 | -18.9601
11.2013 | -34.4676 | -18.9693
11.2074 | -34.468 | -18.9784
11.2135 | -34.4684 | -18.9877
11.2196 | -34.4689 | -18.9969
11.2257 | -34.4692 | -19.0064
11.2318 | -34.4697 | -19.0155
11.2379 | -34.4701 | -19.0248
11.244 | -34.4706 | -19.0341
11.2502 | -34.471 | -19.0432
11.2563 | -34.4714 | -19.0526
11.2624 | -34.4718 | -19.0618
11.2685 | -34.4722 | -19.071
11.2746 | -34.4727 | -19.08
11.2807 | -34.4731 | -19.0892
11.2868 | -34.4735 | -19.0986
11.2929 | -34.474 | -19.1077
11.299 | -34.4744 | -19.117
11.3051 | -34.4749 | -19.1262
11.3112 | -34.4753 | -19.1356
11.3173 | -34.4757 | -19.1445
11.3234 | -34.4762 | -19.1534
11.3295 | -34.4766 | -19.1629
11.3356 | -34.477 | -19.1723
11.3417 | -34.4774 | -19.1814
11.3478 | -34.4779 | -19.1907
11.3539 | -34.4782 | -19.2001
11.36 | -34.4787 | -19.2091
11.3661 | -34.4791 | -19.2181
11.3722 | -34.4795 | -19.2272
11.3783 | -34.48 | -19.2365
11.3844 | -34.4804 | -19.2456
11.3906 | -34.4808 | -19.2548
11.3967 | -34.4812 | -19.264
11.4028 | -34.4817 | -19.2734
11.4089 | -34.4821 | -19.2825
11.415 | -34.4825 | -19.2918
11.4211 | -34.4829 | -19.3012
11.4272 | -34.4833 | -19.3104
11.4333 | -34.4837 | -19.3195
11.4394 | -34.4841 | -19.3286
11.4455 | -34.4845 | -19.3377
11.4516 | -34.4849 | -19.3471
11.4577 | -34.4854 | -19.3562
11.4638 | -34.4858 | -19.3654
11.4699 | -34.4862 | -19.3747
11.476 | -34.4867 | -19.3841
11.4821 | -34.4871 | -19.3932
11.4882 | -34.4875 | -19.4024
11.4943 | -34.4879 | -19.4116
11.5004 | -34.4884 | -19.4205
11.5065 | -34.4888 | -19.4296
11.5126 | -34.4892 | -19.4387
11.5187 | -34.4896 | -19.448
11.5248 | -34.49 | -19.4575
11.5309 | -34.4904 | -19.4666
11.5371 | -34.4909 | -19.4755
11.5432 | -34.4913 | -19.4848
11.5493 | -34.4917 | -19.4943
11.5554 | -34.4922 | -19.5035
11.5615 | -34.4926 | -19.5128
11.5676 | -34.493 | -19.5218
11.5737 | -34.4935 | -19.5309
11.5798 | -34.4939 | -19.54
11.5859 | -34.4943 | -19.5492
11.592 | -34.4947 | -19.5584
11.5981 | -34.4951 | -19.5676
11.6042 | -34.4956 | -19.5768
11.6103 | -34.496 | -19.5858
11.6164 | -34.4964 | -19.5949
11.6225 | -34.4968 | -19.604
11.6286 | -34.4973 | -19.613
11.6347 | -34.4977 | -19.6222
11.6408 | -34.4981 | -19.6313
11.6469 | -34.4985 | -19.6406
11.653 | -34.4989 | -19.6497
11.6591 | -34.4994 | -19.6588
11.6652 | -34.4998 | -19.6678
11.6713 | -34.5002 | -19.6768
11.6775 | -34.5006 | -19.6856
11.6836 | -34.5011 | -19.6946
11.6897 | -34.5015 | -19.7037
11.6958 | -34.5019 | -19.713
11.7019 | -34.5023 | -19.7224
11.708 | -34.5028 | -19.7314
11.7141 | -34.5033 | -19.7402
11.7202 | -34.5037 | -19.7492
11.7263 | -34.5041 | -19.7584
11.7324 | -34.5045 | -19.7675
11.7385 | -34.5049 | -19.7768
11.7446 | -34.5053 | -19.7859
11.7507 | -34.5058 | -19.7949
11.7568 | -34.5062 | -19.804
11.7629 | -34.5066 | -19.8133
11.769 | -34.507 | -19.8226
11.7751 | -34.5075 | -19.8316
11.7812 | -34.5079 | -19.8404
11.7873 | -34.5083 | -19.8495
11.7934 | -34.5088 | -19.8586
11.7995 | -34.5092 | -19.8677
11.8056 | -34.5096 | -19.8768
11.8117 | -34.5101 | -19.886
11.8178 | -34.5105 | -19.8949
11.824 | -34.511 | -19.9038
11.8301 | -34.5114 | -19.9129
11.8362 | -34.5118 | -19.9218
11.8423 | -34.5122 | -19.931
11.8484 | -34.5126 | -19.94
11.8545 | -34.5131 | -19.9489
11.8606 | -34.5135 | -19.9579
11.8667 | -34.514 | -19.967
11.8728 | -34.5144 | -19.976
11.8789 | -34.5148 | -19.9851
11.885 | -34.5153 | -19.9941
11.8911 | -34.5158 | -20.003
11.8972 | -34.5162 | -20.0119
11.9033 | -34.5166 | -20.021
11.9094 | -34.517 | -20.03
11.9155 | -34.5175 | -20.0392
11.9216 | -34.5179 | -20.0483
11.9277 | -34.5183 | -20.0575
11.9338 | -34.5187 | -20.0667
11.9399 | -34.5192 | -20.0759
11.946 | -34.5196 | -20.0848
11.9521 | -34.52 | -20.0937
11.9582 | -34.5204 | -20.1026
11.9644 | -34.5209 | -20.1116
11.9705 | -34.5213 | -20.1206
11.9766 | -34.5218 | -20.1296
11.9827 | -34.5222 | -20.1387
11.9888 | -34.5226 | -20.1477
11.9949 | -34.5231 | -20.1565
12.001 | -34.5235 | -20.1657
12.0071 | -34.5239 | -20.1747
12.0132 | -34.5243 | -20.1837
12.0193 | -34.5248 | -20.1929
12.0254 | -34.5252 | -20.2021
12.0315 | -34.5257 | -20.2107
12.0376 | -34.5261 | -20.2195
12.0437 | -34.5265 | -20.2286
12.0498 | -34.5269 | -20.2376
12.0559 | -34.5274 | -20.2464
12.062 | -34.5278 | -20.2552
12.0681 | -34.5282 | -20.2643
12.0742 | -34.5287 | -20.2731
12.0803 | -34.5291 | -20.2822
12.0864 | -34.5295 | -20.2912
12.0925 | -34.53 | -20.3002
12.0986 | -34.5304 | -20.3091
12.1047 | -34.5308 | -20.3181
12.1109 | -34.5313 | -20.327
12.117 | -34.5317 | -20.3359
12.1231 | -34.5322 | -20.3448
12.1292 | -34.5326 | -20.3538
12.1353 | -34.533 | -20.3626
12.1414 | -34.5334 | -20.3716
12.1475 | -34.5339 | -20.3807
12.1536 | -34.5343 | -20.3897
12.1597 | -34.5347 | -20.3988
12.1658 | -34.5352 | -20.4076
12.1719 | -34.5356 | -20.4164
12.178 | -34.536 | -20.4255
12.1841 | -34.5365 | -20.4347
12.1902 | -34.5369 | -20.4436
12.1963 | -34.5374 | -20.4528
12.2024 | -34.5378 | -20.4616
12.2085 | -34.5383 | -20.4707
12.2146 | -34.5388 | -20.4796
12.2207 | -34.5392 | -20.4885
12.2268 | -34.5396 | -20.4974
12.2329 | -34.5401 | -20.5064
12.239 | -34.5405 | -20.5149
12.2451 | -34.541 | -20.524
12.2513 | -34.5414 | -20.5332
12.2574 | -34.5418 | -20.542
12.2635 | -34.5422 | -20.5509
12.2696 | -34.5427 | -20.5599
12.2757 | -34.5432 | -20.5685
12.2818 | -34.5436 | -20.5773
12.2879 | -34.544 | -20.5864
12.294 | -34.5445 | -20.5952
12.3001 | -34.5449 | -20.604
12.3062 | -34.5454 | -20.6129
12.3123 | -34.5458 | -20.6217
12.3184 | -34.5463 | -20.6303
12.3245 | -34.5467 | -20.6391
12.3306 | -34.5472 | -20.648
12.3367 | -34.5476 | -20.657
12.3428 | -34.548 | -20.6659
12.3489 | -34.5485 | -20.6745
12.355 | -34.549 | -20.6835
12.3611 | -34.5494 | -20.6924
12.3672 | -34.5498 | -20.7011
12.3733 | -34.5502 | -20.7098
12.3794 | -34.5507 | -20.7183
12.3855 | -34.5511 | -20.7269
12.3916 | -34.5516 | -20.7358
12.3978 | -34.5521 | -20.7448
12.4039 | -34.5525 | -20.7539
12.41 | -34.5529 | -20.7629
12.4161 | -34.5534 | -20.7714
12.4222 | -34.5538 | -20.7804
12.4283 | -34.5542 | -20.7892
12.4344 | -34.5547 | -20.7982
12.4405 | -34.5551 | -20.8071
12.4466 | -34.5556 | -20.816
12.4527 | -34.556 | -20.8245
12.4588 | -34.5565 | -20.8333
12.4649 | -34.5569 | -20.8422
12.471 | -34.5574 | -20.8511
12.4771 | -34.5578 | -20.8599
12.4832 | -34.5583 | -20.8687
12.4893 | -34.5587 | -20.8774
12.4954 | -34.5592 | -20.8861
12.5015 | -34.5596 | -20.8949
12.5076 | -34.5601 | -20.9039
12.5137 | -34.5605 | -20.9128
12.5198 | -34.561 | -20.9213
12.5259 | -34.5614 | -20.9303
12.532 | -34.5619 | -20.9392
12.5382 | -34.5623 | -20.948
12.5443 | -34.5627 | -20.9569
12.5504 | -34.5632 | -20.9658
12.5565 | -34.5637 | -20.9743
12.5626 | -34.5642 | -20.9832
12.5687 | -34.5646 | -20.9921
12.5748 | -34.5651 | -21.0009
12.5809 | -34.5655 | -21.0097
12.587 | -34.566 | -21.0185
12.5931 | -34.5665 | -21.0272
12.5992 | -34.5669 | -21.0359
12.6053 | -34.5674 | -21.0447
12.6114 | -34.5678 | -21.0533
12.6175 | -34.5683 | -21.0622
12.6236 | -34.5687 | -21.0708
12.6297 | -34.5692 | -21.0795
12.6358 | -34.5696 | -21.088
12.6419 | -34.5701 | -21.0967
12.648 | -34.5706 | -21.1056
12.6541 | -34.571 | -21.1143
12.6602 | -34.5715 | -21.1228
12.6663 | -34.5719 | -21.1315
12.6724 | -34.5724 | -21.1401
12.6785 | -34.5728 | -21.1487
12.6847 | -34.5733 | -21.1573
12.6908 | -34.5737 | -21.166
12.6969 | -34.5742 | -21.1746
12.703 | -34.5746 | -21.1834
12.7091 | -34.5751 | -21.1921
12.7152 | -34.5755 | -21.2008
12.7213 | -34.576 | -21.2095
12.7274 | -34.5764 | -21.2183
12.7335 | -34.5769 | -21.2268
12.7396 | -34.5773 | -21.2353
12.7457 | -34.5778 | -21.2442
12.7518 | -34.5782 | -21.2528
12.7579 | -34.5787 | -21.2616
12.764 | -34.5792 | -21.2703
12.7701 | -34.5797 | -21.279
12.7762 | -34.5801 | -21.2876
12.7823 | -34.5806 | -21.2963
12.7884 | -34.5811 | -21.305
12.7945 | -34.5815 | -21.3139
12.8006 | -34.582 | -21.3225
12.8067 | -34.5824 | -21.3311
12.8128 | -34.5829 | -21.3396
12.8189 | -34.5834 | -21.3482
12.8251 | -34.5838 | -21.3569
12.8312 | -34.5843 | -21.3658
12.8373 | -34.5848 | -21.3743
12.8434 | -34.5852 | -21.383
12.8495 | -34.5857 | -21.3916
12.8556 | -34.5861 | -21.4002
12.8617 | -34.5866 | -21.4089
12.8678 | -34.587 | -21.4176
12.8739 | -34.5875 | -21.4262
12.88 | -34.5879 | -21.4347
12.8861 | -34.5884 | -21.4434
12.8922 | -34.5888 | -21.4521
12.8983 | -34.5893 | -21.4608
12.9044 | -34.5898 | -21.4694
12.9105 | -34.5903 | -21.4777
12.9166 | -34.5907 | -21.4863
12.9227 | -34.5912 | -21.4949
12.9288 | -34.5917 | -21.5036
12.9349 | -34.5921 | -21.5123
12.941 | -34.5926 | -21.5208
12.9471 | -34.5931 | -21.5295
12.9532 | -34.5935 | -21.5381
12.9593 | -34.594 | -21.5469
12.9654 | -34.5945 | -21.5556
12.9716 | -34.5949 | -21.5643
12.9777 | -34.5954 | -21.5729
12.9838 | -34.5959 | -21.5814
12.9899 | -34.5963 | -21.59
12.996 | -34.5968 | -21.5987
13.0021 | -34.5972 | -21.6074
13.0082 | -34.5977 | -21.6159
13.0143 | -34.5982 | -21.6242
13.0204 | -34.5987 | -21.6326
13.0265 | -34.5991 | -21.641
13.0326 | -34.5996 | -21.6496
13.0387 | -34.6001 | -21.6584
13.0448 | -34.6005 | -21.6669
13.0509 | -34.601 | -21.6754
13.057 | -34.6015 | -21.6838
13.0631 | -34.602 | -21.6924
13.0692 | -34.6024 | -21.7009
13.0753 | -34.6029 | -21.7095
13.0814 | -34.6034 | -21.7183
13.0875 | -34.6039 | -21.7268
13.0936 | -34.6043 | -21.7354
13.0997 | -34.6048 | -21.7438
13.1058 | -34.6053 | -21.7523
13.112 | -34.6057 | -21.7608
13.1181 | -34.6062 | -21.7692
13.1242 | -34.6067 | -21.7779
13.1303 | -34.6072 | -21.7865
13.1364 | -34.6076 | -21.7952
13.1425 | -34.6081 | -21.8036
13.1486 | -34.6085 | -21.8122
13.1547 | -34.609 | -21.8206
13.1608 | -34.6095 | -21.8291
13.1669 | -34.6099 | -21.8375
13.173 | -34.6104 | -21.8462
13.1791 | -34.6109 | -21.8547
13.1852 | -34.6113 | -21.8633
13.1913 | -34.6118 | -21.8717
13.1974 | -34.6123 | -21.8803
13.2035 | -34.6127 | -21.8889
13.2096 | -34.6132 | -21.8973
13.2157 | -34.6137 | -21.9059
13.2218 | -34.6141 | -21.9144
13.2279 | -34.6146 | -21.9228
13.234 | -34.6151 | -21.9313
13.2401 | -34.6156 | -21.9398
13.2462 | -34.616 | -21.9482
13.2524 | -34.6165 | -21.9567
13.2585 | -34.617 | -21.9651
13.2646 | -34.6174 | -21.9736
13.2707 | -34.6179 | -21.9821
13.2768 | -34.6184 | -21.9907
13.2829 | -34.6189 | -21.9993
13.289 | -34.6193 | -22.0077
13.2951 | -34.6198 | -22.016
13.3012 | -34.6203 | -22.0246
13.3073 | -34.6208 | -22.0331
13.3134 | -34.6212 | -22.0417
13.3195 | -34.6217 | -22.0501
13.3256 | -34.6222 | -22.0586
13.3317 | -34.6226 | -22.0669
13.3378 | -34.6231 | -22.0754
13.3439 | -34.6236 | -22.084
13.35 | -34.6241 | -22.0925
13.3561 | -34.6245 | -22.1009
13.3622 | -34.625 | -22.1094
13.3683 | -34.6255 | -22.1178
13.3744 | -34.6259 | -22.1261
13.3805 | -34.6264 | -22.1347
13.3866 | -34.6269 | -22.143
13.3927 | -34.6273 | -22.1516
13.3989 | -34.6278 | -22.16
13.405 | -34.6283 | -22.1684
13.4111 | -34.6287 | -22.1767
13.4172 | -34.6292 | -22.1852
13.4233 | -34.6297 | -22.1937
13.4294 | -34.6301 | -22.2023
13.4355 | -34.6306 | -22.2104
13.4416 | -34.6311 | -22.2188
13.4477 | -34.6316 | -22.2273
13.4538 | -34.632 | -22.2357
13.4599 | -34.6325 | -22.2441
13.466 | -34.633 | -22.2526
13.4721 | -34.6335 | -22.2609
13.4782 | -34.634 | -22.2694
13.4843 | -34.6344 | -22.2779
13.4904 | -34.6349 | -22.2863
13.4965 | -34.6354 | -22.2949
13.5026 | -34.6359 | -22.3034
13.5087 | -34.6363 | -22.3116
13.5148 | -34.6368 | -22.3201
13.5209 | -34.6373 | -22.3287
13.527 | -34.6377 | -22.337
13.5331 | -34.6382 | -22.3454
13.5393 | -34.6387 | -22.3536
13.5454 | -34.6392 | -22.3621
13.5515 | -34.6397 | -22.3707
13.5576 | -34.6402 | -22.3791
13.5637 | -34.6406 | -22.3876
13.5698 | -34.6411 | -22.396
13.5759 | -34.6416 | -22.4042
13.582 | -34.642 | -22.4126
13.5881 | -34.6425 | -22.421
13.5942 | -34.643 | -22.4295
13.6003 | -34.6435 | -22.4379
13.6064 | -34.6439 | -22.4466
13.6125 | -34.6444 | -22.455
13.6186 | -34.6449 | -22.4633
13.6247 | -34.6454 | -22.4718
13.6308 | -34.6459 | -22.4803
13.6369 | -34.6464 | -22.4888
13.643 | -34.6468 | -22.4974
13.6491 | -34.6473 | -22.5057
13.6552 | -34.6478 | -22.5141
13.6613 | -34.6483 | -22.5224
13.6674 | -34.6487 | -22.5309
13.6735 | -34.6492 | -22.5393
13.6796 | -34.6497 | -22.5476
13.6858 | -34.6502 | -22.5559
13.6919 | -34.6507 | -22.5642
13.698 | -34.6512 | -22.5726
13.7041 | -34.6516 | -22.5811
13.7102 | -34.6521 | -22.5893
13.7163 | -34.6526 | -22.5977
13.7224 | -34.6531 | -22.6061
13.7285 | -34.6536 | -22.6146
13.7346 | -34.654 | -22.6229
13.7407 | -34.6545 | -22.6311
13.7468 | -34.655 | -22.6397
13.7529 | -34.6554 | -22.6479
13.759 | -34.6559 | -22.6564
13.7651 | -34.6564 | -22.6648
13.7712 | -34.6569 | -22.6732
13.7773 | -34.6573 | -22.6817
13.7834 | -34.6578 | -22.6901
13.7895 | -34.6583 | -22.6983
13.7956 | -34.6588 | -22.7068
13.8017 | -34.6593 | -22.7151
13.8078 | -34.6597 | -22.7235
13.8139 | -34.6602 | -22.7318
13.82 | -34.6607 | -22.7401
13.8262 | -34.6611 | -22.7484
13.8323 | -34.6616 | -22.7568
13.8384 | -34.6621 | -22.7652
13.8445 | -34.6626 | -22.7735
13.8506 | -34.6631 | -22.7819
13.8567 | -34.6636 | -22.7902
13.8628 | -34.664 | -22.7985
13.8689 | -34.6645 | -22.8071
13.875 | -34.665 | -22.8154
13.8811 | -34.6654 | -22.8237
13.8872 | -34.6659 | -22.832
13.8933 | -34.6664 | -22.8402
13.8994 | -34.6669 | -22.8485
13.9055 | -34.6673 | -22.857
13.9116 | -34.6678 | -22.8652
13.9177 | -34.6683 | -22.8735
13.9238 | -34.6688 | -22.8819
13.9299 | -34.6693 | -22.8903
13.936 | -34.6697 | -22.8986
13.9421 | -34.6702 | -22.9069
13.9482 | -34.6707 | -22.9152
13.9543 | -34.6711 | -22.9238
13.9604 | -34.6716 | -22.9321
13.9665 | -34.6721 | -22.9404
13.9727 | -34.6726 | -22.9489
13.9788 | -34.6731 | -22.9572
13.9849 | -34.6736 | -22.9656
13.991 | -34.674 | -22.9739
13.9971 | -34.6745 | -22.9822
14.0032 | -34.675 | -22.9905
14.0093 | -34.6755 | -22.999
14.0154 | -34.6759 | -23.0073
14.0215 | -34.6764 | -23.0156
14.0276 | -34.6769 | -23.0238
14.0337 | -34.6774 | -23.032
14.0398 | -34.6779 | -23.0403
14.0459 | -34.6784 | -23.0487
14.052 | -34.6788 | -23.0571
14.0581 | -34.6793 | -23.0655
14.0642 | -34.6797 | -23.0738
14.0703 | -34.6802 | -23.0821
14.0764 | -34.6807 | -23.0903
14.0825 | -34.6812 | -23.0987
14.0886 | -34.6817 | -23.1071
14.0947 | -34.6822 | -23.1154
14.1008 | -34.6826 | -23.1237
14.1069 | -34.6831 | -23.132
14.1131 | -34.6836 | -23.1403
14.1192 | -34.684 | -23.1487
14.1253 | -34.6845 | -23.1569
14.1314 | -34.685 | -23.1652
14.1375 | -34.6855 | -23.1735
14.1436 | -34.686 | -23.1817
14.1497 | -34.6864 | -23.19
14.1558 | -34.6869 | -23.1984
14.1619 | -34.6874 | -23.2067
14.168 | -34.6879 | -23.215
14.1741 | -34.6884 | -23.2231
14.1802 | -34.6889 | -23.2313
14.1863 | -34.6894 | -23.2398
14.1924 | -34.6898 | -23.2481
14.1985 | -34.6903 | -23.2565
14.2046 | -34.6908 | -23.2649
14.2107 | -34.6913 | -23.2729
14.2168 | -34.6918 | -23.2813
14.2229 | -34.6922 | -23.2896
14.229 | -34.6927 | -23.2979
14.2351 | -34.6932 | -23.3062
14.2412 | -34.6937 | -23.3146
14.2473 | -34.6942 | -23.3228
14.2534 | -34.6947 | -23.331
14.2596 | -34.6951 | -23.3393
14.2657 | -34.6956 | -23.3475
14.2718 | -34.6961 | -23.3559
14.2779 | -34.6966 | -23.3642
14.284 | -34.6971 | -23.3724
14.2901 | -34.6975 | -23.3806
14.2962 | -34.698 | -23.3889
14.3023 | -34.6985 | -23.3973
14.3084 | -34.699 | -23.4055
14.3145 | -34.6995 | -23.4137
14.3206 | -34.7 | -23.4221
14.3267 | -34.7005 | -23.4304
14.3328 | -34.7009 | -23.4386
14.3389 | -34.7014 | -23.4468
14.345 | -34.7019 | -23.4551
14.3511 | -34.7024 | -23.4633
14.3572 | -34.7029 | -23.4717
14.3633 | -34.7033 | -23.4801
14.3694 | -34.7038 | -23.4882
14.3755 | -34.7043 | -23.4966
14.3816 | -34.7048 | -23.505
14.3877 | -34.7053 | -23.5132
14.3938 | -34.7058 | -23.5217
14.4 | -34.7063 | -23.5298
14.4061 | -34.7067 | -23.5381
14.4122 | -34.7072 | -23.5463
14.4183 | -34.7077 | -23.5545
14.4244 | -34.7082 | -23.5626
14.4305 | -34.7087 | -23.5709
14.4366 | -34.7092 | -23.5791
14.4427 | -34.7097 | -23.5875
14.4488 | -34.7101 | -23.5958
14.4549 | -34.7106 | -23.6039
14.461 | -34.7111 | -23.6121
14.4671 | -34.7116 | -23.6205
14.4732 | -34.7121 | -23.6288
14.4793 | -34.7125 | -23.637
14.4854 | -34.713 | -23.6452
14.4915 | -34.7135 | -23.6531
14.4976 | -34.714 | -23.6612
14.5037 | -34.7145 | -23.6694
14.5098 | -34.715 | -23.6777
14.5159 | -34.7155 | -23.686
14.522 | -34.7159 | -23.6941
14.5281 | -34.7164 | -23.7024
14.5342 | -34.7169 | -23.7107
14.5403 | -34.7174 | -23.719
14.5465 | -34.7179 | -23.7273
14.5526 | -34.7183 | -23.7357
14.5587 | -34.7188 | -23.7439
14.5648 | -34.7193 | -23.7521
14.5709 | -34.7198 | -23.7602
14.577 | -34.7203 | -23.7684
14.5831 | -34.7207 | -23.7767
14.5892 | -34.7212 | -23.7849
14.5953 | -34.7217 | -23.7928
14.6014 | -34.7222 | -23.801
14.6075 | -34.7227 | -23.8092
14.6136 | -34.7231 | -23.8174
14.6197 | -34.7236 | -23.8256
14.6258 | -34.7241 | -23.8339
14.6319 | -34.7246 | -23.8422
14.638 | -34.725 | -23.8504
14.6441 | -34.7255 | -23.8587
14.6502 | -34.726 | -23.8671
14.6563 | -34.7265 | -23.8752
14.6624 | -34.727 | -23.8834
14.6685 | -34.7275 | -23.8913
14.6746 | -34.7279 | -23.8995
14.6807 | -34.7284 | -23.9077
14.6869 | -34.7289 | -23.9159
14.693 | -34.7294 | -23.9241
14.6991 | -34.7299 | -23.9323
14.7052 | -34.7304 | -23.9403
14.7113 | -34.7309 | -23.9487
14.7174 | -34.7314 | -23.957
14.7235 | -34.7319 | -23.9651
14.7296 | -34.7323 | -23.9733
14.7357 | -34.7328 | -23.9816
14.7418 | -34.7333 | -23.9897
14.7479 | -34.7338 | -23.9979
14.754 | -34.7343 | -24.0062
14.7601 | -34.7348 | -24.0143
14.7662 | -34.7353 | -24.0224
14.7723 | -34.7357 | -24.0304
14.7784 | -34.7362 | -24.0386
14.7845 | -34.7367 | -24.0468
14.7906 | -34.7372 | -24.0552
14.7967 | -34.7377 | -24.0634
14.8028 | -34.7382 | -24.0716
14.8089 | -34.7386 | -24.0796
14.815 | -34.7391 | -24.0878
14.8211 | -34.7396 | -24.0961
14.8272 | -34.7401 | -24.1044
14.8334 | -34.7406 | -24.1127
14.8395 | -34.7411 | -24.1208
14.8456 | -34.7415 | -24.129
14.8517 | -34.742 | -24.1372
14.8578 | -34.7425 | -24.1456
14.8639 | -34.743 | -24.1537
14.87 | -34.7435 | -24.162
14.8761 | -34.744 | -24.1702
14.8822 | -34.7445 | -24.1784
14.8883 | -34.745 | -24.1867
14.8944 | -34.7454 | -24.1948
14.9005 | -34.7459 | -24.2031
14.9066 | -34.7464 | -24.2113
14.9127 | -34.7469 | -24.2195
14.9188 | -34.7474 | -24.2276
14.9249 | -34.7479 | -24.236
14.931 | -34.7484 | -24.2441
14.9371 | -34.7489 | -24.2522
14.9432 | -34.7493 | -24.2604
14.9493 | -34.7498 | -24.2685
14.9554 | -34.7503 | -24.2767
14.9615 | -34.7508 | -24.2848
14.9676 | -34.7513 | -24.2932
14.9738 | -34.7518 | -24.3013
14.9799 | -34.7523 | -24.3096
14.986 | -34.7528 | -24.3176
14.9921 | -34.7533 | -24.3257
14.9982 | -34.7537 | -24.3338
15.0043 | -34.7542 | -24.3419
15.0104 | -34.7547 | -24.3501
15.0165 | -34.7552 | -24.3584
15.0226 | -34.7557 | -24.3665
15.0287 | -34.7562 | -24.3746
15.0348 | -34.7567 | -24.3828
15.0409 | -34.7571 | -24.3909
15.047 | -34.7577 | -24.3991
15.0531 | -34.7582 | -24.407
15.0592 | -34.7587 | -24.4152
15.0653 | -34.7591 | -24.4234
15.0714 | -34.7596 | -24.4316
15.0775 | -34.7601 | -24.4396
15.0836 | -34.7606 | -24.4477
15.0897 | -34.7611 | -24.4556
15.0958 | -34.7616 | -24.4638
15.1019 | -34.7621 | -24.4721
15.108 | -34.7626 | -24.4802
15.1141 | -34.7631 | -24.4883
15.1203 | -34.7636 | -24.4963
15.1264 | -34.7641 | -24.5045
15.1325 | -34.7646 | -24.5126
15.1386 | -34.7651 | -24.5207
15.1447 | -34.7656 | -24.5288
15.1508 | -34.7661 | -24.537
15.1569 | -34.7665 | -24.545
15.163 | -34.767 | -24.5529
15.1691 | -34.7675 | -24.5609
15.1752 | -34.768 | -24.569
15.1813 | -34.7685 | -24.5772
15.1874 | -34.769 | -24.5853
15.1935 | -34.7695 | -24.5932
15.1996 | -34.77 | -24.6015
15.2057 | -34.7705 | -24.6096
15.2118 | -34.771 | -24.6177
15.2179 | -34.7715 | -24.626
15.224 | -34.772 | -24.6341
15.2301 | -34.7725 | -24.642
15.2362 | -34.7729 | -24.6501
15.2423 | -34.7734 | -24.6583
15.2484 | -34.7739 | -24.6665
15.2545 | -34.7744 | -24.6749
15.2607 | -34.7749 | -24.683
15.2668 | -34.7754 | -24.691
15.2729 | -34.7759 | -24.6991
15.279 | -34.7764 | -24.7074
15.2851 | -34.7769 | -24.7156
15.2912 | -34.7774 | -24.7238
15.2973 | -34.7779 | -24.7321
15.3034 | -34.7784 | -24.7402
15.3095 | -34.7789 | -24.7483
15.3156 | -34.7794 | -24.7564
15.3217 | -34.7799 | -24.7644
15.3278 | -34.7804 | -24.7727
15.3339 | -34.7809 | -24.7807
15.34 | -34.7814 | -24.7886
15.3461 | -34.7819 | -24.7966
15.3522 | -34.7824 | -24.8047
15.3583 | -34.7829 | -24.8128
15.3644 | -34.7833 | -24.8208
15.3705 | -34.7838 | -24.8287
15.3766 | -34.7843 | -24.837
15.3827 | -34.7848 | -24.845
15.3888 | -34.7853 | -24.853
15.3949 | -34.7858 | -24.8612
15.401 | -34.7863 | -24.8695
15.4072 | -34.7868 | -24.8774
15.4133 | -34.7873 | -24.8856
15.4194 | -34.7878 | -24.8936
15.4255 | -34.7883 | -24.9017
15.4316 | -34.7888 | -24.9098
15.4377 | -34.7892 | -24.9179
15.4438 | -34.7898 | -24.9259
15.4499 | -34.7903 | -24.9341
15.456 | -34.7908 | -24.9421
15.4621 | -34.7913 | -24.9504
15.4682 | -34.7918 | -24.9586
15.4743 | -34.7923 | -24.9666
15.4804 | -34.7928 | -24.9746
15.4865 | -34.7933 | -24.9826
15.4926 | -34.7938 | -24.9907
15.4987 | -34.7943 | -24.999
15.5048 | -34.7948 | -25.0071
15.5109 | -34.7953 | -25.0153
15.517 | -34.7958 | -25.0235
15.5231 | -34.7963 | -25.0317
15.5292 | -34.7967 | -25.0398
15.5353 | -34.7972 | -25.048
15.5414 | -34.7977 | -25.056
15.5476 | -34.7982 | -25.0637
15.5537 | -34.7987 | -25.0717
15.5598 | -34.7992 | -25.0799
15.5659 | -34.7997 | -25.088
15.572 | -34.8002 | -25.0961
15.5781 | -34.8007 | -25.1042
15.5842 | -34.8012 | -25.1121
15.5903 | -34.8017 | -25.12
15.5964 | -34.8022 | -25.1281
15.6025 | -34.8027 | -25.1361
15.6086 | -34.8031 | -25.1441
15.6147 | -34.8036 | -25.1522
15.6208 | -34.8041 | -25.1601
15.6269 | -34.8046 | -25.1681
15.633 | -34.805 | -25.1764
15.6391 | -34.8055 | -25.1844
15.6452 | -34.806 | -25.1927
15.6513 | -34.8065 | -25.2005
15.6574 | -34.807 | -25.2087
15.6635 | -34.8076 | -25.2168
15.6696 | -34.8081 | -25.225
15.6757 | -34.8086 | -25.233
15.6818 | -34.8091 | -25.2409
15.688 | -34.8096 | -25.2488
15.6941 | -34.8101 | -25.2567
15.7002 | -34.8105 | -25.2647
15.7063 | -34.811 | -25.2727
15.7124 | -34.8115 | -25.2808
15.7185 | -34.812 | -25.2888
15.7246 | -34.8125 | -25.2967
15.7307 | -34.813 | -25.3046
15.7368 | -34.8135 | -25.3127
15.7429 | -34.814 | -25.3206
15.749 | -34.8145 | -25.3287
15.7551 | -34.815 | -25.3367
15.7612 | -34.8155 | -25.3447
15.7673 | -34.816 | -25.3526
15.7734 | -34.8165 | -25.3605
15.7795 | -34.8169 | -25.3685
15.7856 | -34.8174 | -25.3762
15.7917 | -34.818 | -25.3839
15.7978 | -34.8185 | -25.3919
15.8039 | -34.819 | -25.4
15.81 | -34.8195 | -25.408
15.8161 | -34.82 | -25.4159
15.8222 | -34.8205 | -25.4239
15.8283 | -34.821 | -25.4316
15.8345 | -34.8215 | -25.4397
15.8406 | -34.822 | -25.4476
15.8467 | -34.8225 | -25.4558
15.8528 | -34.823 | -25.4639
15.8589 | -34.8235 | -25.4719
15.865 | -34.824 | -25.4796
15.8711 | -34.8245 | -25.4877
15.8772 | -34.825 | -25.4954
15.8833 | -34.8255 | -25.5035
15.8894 | -34.826 | -25.5116
15.8955 | -34.8265 | -25.5195
15.9016 | -34.827 | -25.5274
15.9077 | -34.8274 | -25.5353
15.9138 | -34.8279 | -25.5434
15.9199 | -34.8284 | -25.5514
15.926 | -34.8289 | -25.5594
15.9321 | -34.8294 | -25.5674
15.9382 | -34.8299 | -25.5754
15.9443 | -34.8304 | -25.5834
15.9504 | -34.8309 | -25.5914
15.9565 | -34.8314 | -25.5993
15.9626 | -34.8319 | -25.6073
15.9687 | -34.8324 | -25.6149
15.9749 | -34.8329 | -25.6228
15.981 | -34.8334 | -25.6308
15.9871 | -34.8339 | -25.639
15.9932 | -34.8344 | -25.6469
15.9993 | -34.8349 | -25.6548
16.0054 | -34.8355 | -25.6624
16.0115 | -34.836 | -25.6703
16.0176 | -34.8365 | -25.6782
16.0237 | -34.837 | -25.6862
16.0298 | -34.8374 | -25.6942
16.0359 | -34.8379 | -25.7022
16.042 | -34.8385 | -25.7099
16.0481 | -34.8389 | -25.7177
16.0542 | -34.8395 | -25.7258
16.0603 | -34.84 | -25.7337
16.0664 | -34.8404 | -25.7417
16.0725 | -34.8409 | -25.7494
16.0786 | -34.8414 | -25.7574
16.0847 | -34.8419 | -25.7655
16.0908 | -34.8424 | -25.7735
16.0969 | -34.8429 | -25.7815
16.103 | -34.8434 | -25.7894
16.1091 | -34.8439 | -25.7971
16.1152 | -34.8444 | -25.8049
16.1214 | -34.8449 | -25.8129
16.1275 | -34.8454 | -25.8208
16.1336 | -34.8459 | -25.8285
16.1397 | -34.8464 | -25.8364
16.1458 | -34.8469 | -25.8442
16.1519 | -34.8474 | -25.8522
16.158 | -34.8479 | -25.8603
16.1641 | -34.8484 | -25.8683
16.1702 | -34.8489 | -25.8763
16.1763 | -34.8494 | -25.8843
16.1824 | -34.8499 | -25.8921
16.1885 | -34.8504 | -25.9
16.1946 | -34.8509 | -25.9077
16.2007 | -34.8514 | -25.9158
16.2068 | -34.8519 | -25.9235
16.2129 | -34.8524 | -25.9315
16.219 | -34.8529 | -25.9392
16.2251 | -34.8534 | -25.9468
16.2312 | -34.8539 | -25.9547
16.2373 | -34.8545 | -25.9626
16.2434 | -34.855 | -25.9705
16.2495 | -34.8555 | -25.9783
16.2556 | -34.856 | -25.9863
16.2618 | -34.8565 | -25.9942
16.2679 | -34.857 | -26.0022
16.274 | -34.8575 | -26.01
16.2801 | -34.858 | -26.0179
16.2862 | -34.8585 | -26.0257
16.2923 | -34.859 | -26.0337
16.2984 | -34.8595 | -26.0415
16.3045 | -34.86 | -26.0494
16.3106 | -34.8605 | -26.0573
16.3167 | -34.861 | -26.0652
16.3228 | -34.8615 | -26.0729
16.3289 | -34.862 | -26.0806
16.335 | -34.8625 | -26.0886
16.3411 | -34.863 | -26.0965
16.3472 | -34.8635 | -26.1045
16.3533 | -34.864 | -26.1124
16.3594 | -34.8645 | -26.1202
16.3655 | -34.865 | -26.1278
16.3716 | -34.8655 | -26.1356
16.3777 | -34.866 | -26.1435
16.3838 | -34.8665 | -26.1514
16.3899 | -34.867 | -26.1594
16.396 | -34.8675 | -26.1675
16.4021 | -34.868 | -26.1752
16.4083 | -34.8685 | -26.1829
16.4144 | -34.869 | -26.1907
16.4205 | -34.8695 | -26.1985
16.4266 | -34.87 | -26.2063
16.4327 | -34.8705 | -26.2142
16.4388 | -34.871 | -26.2221
16.4449 | -34.8715 | -26.2299
16.451 | -34.8719 | -26.2379
16.4571 | -34.8725 | -26.2458
16.4632 | -34.8729 | -26.2533
16.4693 | -34.8734 | -26.2611
16.4754 | -34.874 | -26.269
16.4815 | -34.8745 | -26.2767
16.4876 | -34.875 | -26.2843
16.4937 | -34.8754 | -26.2921
16.4998 | -34.876 | -26.2995
16.5059 | -34.8765 | -26.3075
16.512 | -34.877 | -26.3154
16.5181 | -34.8775 | -26.3235
16.5242 | -34.878 | -26.3314
16.5303 | -34.8785 | -26.3392
16.5364 | -34.879 | -26.347
16.5425 | -34.8795 | -26.355
16.5487 | -34.8799 | -26.3628
16.5548 | -34.8805 | -26.3708
16.5609 | -34.8809 | -26.3786
16.567 | -34.8815 | -26.3862
16.5731 | -34.882 | -26.3943
16.5792 | -34.8825 | -26.4023
16.5853 | -34.883 | -26.4102
16.5914 | -34.8835 | -26.4182
16.5975 | -34.884 | -26.4262
16.6036 | -34.8845 | -26.4341
16.6097 | -34.885 | -26.4418
16.6158 | -34.8855 | -26.4497
16.6219 | -34.886 | -26.4573
16.628 | -34.8865 | -26.4651
16.6341 | -34.887 | -26.4731
16.6402 | -34.8875 | -26.4808
16.6463 | -34.8879 | -26.4886
16.6524 | -34.8884 | -26.4963
16.6585 | -34.8889 | -26.5043
16.6646 | -34.8894 | -26.512
16.6707 | -34.89 | -26.5196
16.6768 | -34.8904 | -26.5275
16.6829 | -34.8909 | -26.5352
16.689 | -34.8914 | -26.543
16.6952 | -34.8919 | -26.5506
16.7013 | -34.8924 | -26.5584
16.7074 | -34.8928 | -26.5661
16.7135 | -34.8934 | -26.574
16.7196 | -34.8939 | -26.5817
16.7257 | -34.8944 | -26.5893
16.7318 | -34.8949 | -26.5971
16.7379 | -34.8954 | -26.6049
16.744 | -34.8959 | -26.6125
16.7501 | -34.8964 | -26.6206
16.7562 | -34.8969 | -26.6285
16.7623 | -34.8974 | -26.6361
16.7684 | -34.8979 | -26.6438
16.7745 | -34.8984 | -26.6517
16.7806 | -34.8989 | -26.6592
16.7867 | -34.8994 | -26.6669
16.7928 | -34.8999 | -26.6747
16.7989 | -34.9004 | -26.6825
16.805 | -34.9009 | -26.6904
16.8111 | -34.9014 | -26.6979
16.8172 | -34.9019 | -26.7058
16.8233 | -34.9024 | -26.7137
16.8294 | -34.9029 | -26.7214
16.8356 | -34.9034 | -26.7294
16.8417 | -34.9039 | -26.7373
16.8478 | -34.9044 | -26.745
16.8539 | -34.9049 | -26.7527
16.86 | -34.9054 | -26.7606
16.8661 | -34.9059 | -26.7685
16.8722 | -34.9064 | -26.7763
16.8783 | -34.9068 | -26.784
16.8844 | -34.9074 | -26.7916
16.8905 | -34.9079 | -26.7995
16.8966 | -34.9084 | -26.8073
16.9027 | -34.9089 | -26.8151
16.9088 | -34.9094 | -26.8231
16.9149 | -34.9099 | -26.8308
16.921 | -34.9104 | -26.8383
16.9271 | -34.9109 | -26.8462
16.9332 | -34.9114 | -26.8541
16.9393 | -34.9119 | -26.8617
16.9454 | -34.9124 | -26.8692
16.9515 | -34.9129 | -26.8771
16.9576 | -34.9134 | -26.8848
16.9637 | -34.9139 | -26.8924
16.9698 | -34.9144 | -26.9001
16.9759 | -34.9149 | -26.9079
16.9821 | -34.9154 | -26.9159
16.9882 | -34.9159 | -26.9236
16.9943 | -34.9164 | -26.9315
17.0004 | -34.9169 | -26.9392
17.0065 | -34.9174 | -26.947
17.0126 | -34.9179 | -26.9547
17.0187 | -34.9184 | -26.9624
17.0248 | -34.9189 | -26.9701
17.0309 | -34.9194 | -26.9776
17.037 | -34.9199 | -26.9852
17.0431 | -34.9204 | -26.9931
17.0492 | -34.9209 | -27.0009
17.0553 | -34.9214 | -27.0085
17.0614 | -34.9219 | -27.016
17.0675 | -34.9224 | -27.0235
17.0736 | -34.9229 | -27.0313
17.0797 | -34.9234 | -27.0389
17.0858 | -34.9239 | -27.0466
17.0919 | -34.9244 | -27.0543
17.098 | -34.9249 | -27.0621
17.1041 | -34.9253 | -27.0699
17.1102 | -34.9259 | -27.0777
17.1163 | -34.9264 | -27.0853
17.1225 | -34.9269 | -27.093
17.1286 | -34.9274 | -27.1004
17.1347 | -34.9279 | -27.1083
17.1408 | -34.9284 | -27.1161
17.1469 | -34.929 | -27.1237
17.153 | -34.9295 | -27.1313
17.1591 | -34.93 | -27.1391
17.1652 | -34.9305 | -27.1466
17.1713 | -34.931 | -27.1543
17.1774 | -34.9315 | -27.162
17.1835 | -34.932 | -27.1696
17.1896 | -34.9325 | -27.1774
17.1957 | -34.933 | -27.1849
17.2018 | -34.9335 | -27.1922
17.2079 | -34.934 | -27.1998
17.214 | -34.9345 | -27.2075
17.2201 | -34.935 | -27.2152
17.2262 | -34.9355 | -27.223
17.2323 | -34.936 | -27.2306
17.2384 | -34.9365 | -27.2382
17.2445 | -34.937 | -27.2458
17.2506 | -34.9375 | -27.2533
17.2567 | -34.938 | -27.2611
17.2628 | -34.9385 | -27.2688
17.269 | -34.939 | -27.2762
17.2751 | -34.9395 | -27.2838
17.2812 | -34.94 | -27.2915
17.2873 | -34.9405 | -27.2993
17.2934 | -34.941 | -27.307
17.2995 | -34.9415 | -27.3148
17.3056 | -34.942 | -27.3226
17.3117 | -34.9425 | -27.3303
17.3178 | -34.943 | -27.3379
17.3239 | -34.9435 | -27.3457
17.33 | -34.944 | -27.3532
17.3361 | -34.9445 | -27.3607
17.3422 | -34.945 | -27.3684
17.3483 | -34.9454 | -27.3761
17.3544 | -34.9459 | -27.3837
17.3605 | -34.9464 | -27.3913
17.3666 | -34.9469 | -27.399
17.3727 | -34.9474 | -27.4067
17.3788 | -34.9479 | -27.4142
17.3849 | -34.9484 | -27.4219
17.391 | -34.9488 | -27.4296
17.3971 | -34.9493 | -27.4373
17.4032 | -34.9498 | -27.4448
17.4094 | -34.9503 | -27.4522
17.4155 | -34.9508 | -27.4597
17.4216 | -34.9513 | -27.4673
17.4277 | -34.9518 | -27.475
17.4338 | -34.9523 | -27.4826
17.4399 | -34.9528 | -27.4905
17.446 | -34.9533 | -27.4979
17.4521 | -34.9538 | -27.5056
17.4582 | -34.9543 | -27.5131
17.4643 | -34.9548 | -27.5207
17.4704 | -34.9553 | -27.5283
17.4765 | -34.9558 | -27.5361
17.4826 | -34.9563 | -27.5438
17.4887 | -34.9567 | -27.5516
17.4948 | -34.9572 | -27.559
17.5009 | -34.9577 | -27.5666
17.507 | -34.9582 | -27.5745
17.5131 | -34.9587 | -27.582
17.5192 | -34.9592 | -27.5894
17.5253 | -34.9597 | -27.597
17.5314 | -34.9602 | -27.6045
17.5375 | -34.9606 | -27.6121
17.5436 | -34.9611 | -27.6199
17.5497 | -34.9616 | -27.6275
17.5559 | -34.9621 | -27.6349
17.562 | -34.9626 | -27.6425
17.5681 | -34.9631 | -27.6502
17.5742 | -34.9636 | -27.6579
17.5803 | -34.9641 | -27.6656
17.5864 | -34.9647 | -27.6729
17.5925 | -34.9651 | -27.6808
17.5986 | -34.9656 | -27.6882
17.6047 | -34.9661 | -27.6957
17.6108 | -34.9666 | -27.7032
17.6169 | -34.9671 | -27.7108
17.623 | -34.9676 | -27.7184
17.6291 | -34.9681 | -27.7257
17.6352 | -34.9686 | -27.7331
17.6413 | -34.9691 | -27.7409
17.6474 | -34.9696 | -27.7484
17.6535 | -34.9701 | -27.7557
17.6596 | -34.9707 | -27.7631
17.6657 | -34.9712 | -27.7705
17.6718 | -34.9716 | -27.778
17.6779 | -34.9721 | -27.7855
17.684 | -34.9726 | -27.7929
17.6901 | -34.9731 | -27.8006
17.6963 | -34.9736 | -27.808
17.7024 | -34.9741 | -27.8156
17.7085 | -34.9746 | -27.8231
17.7146 | -34.9751 | -27.8306
17.7207 | -34.9755 | -27.8381
17.7268 | -34.976 | -27.8455
17.7329 | -34.9765 | -27.8531
17.739 | -34.977 | -27.8609
17.7451 | -34.9775 | -27.8684
17.7512 | -34.978 | -27.876
17.7573 | -34.9785 | -27.8837
17.7634 | -34.979 | -27.891
17.7695 | -34.9795 | -27.8986
17.7756 | -34.98 | -27.9061
17.7817 | -34.9805 | -27.9136
17.7878 | -34.981 | -27.9213
17.7939 | -34.9815 | -27.9289
17.8 | -34.982 | -27.9364
17.8061 | -34.9824 | -27.9439
17.8122 | -34.9829 | -27.9514
17.8183 | -34.9834 | -27.9588
17.8244 | -34.9839 | -27.9662
17.8305 | -34.9844 | -27.9736
17.8366 | -34.9849 | -27.9808
17.8428 | -34.9854 | -27.9885
17.8489 | -34.9859 | -27.9961
17.855 | -34.9864 | -28.0034
17.8611 | -34.9869 | -28.0108
17.8672 | -34.9873 | -28.0182
17.8733 | -34.9878 | -28.0259
17.8794 | -34.9883 | -28.0333
17.8855 | -34.9888 | -28.0409
17.8916 | -34.9893 | -28.0484
17.8977 | -34.9898 | -28.056
17.9038 | -34.9903 | -28.0635
17.9099 | -34.9907 | -28.0712
17.916 | -34.9912 | -28.0788
17.9221 | -34.9917 | -28.0864
17.9282 | -34.9922 | -28.0939
17.9343 | -34.9927 | -28.1017
17.9404 | -34.9932 | -28.109
17.9465 | -34.9937 | -28.1166
17.9526 | -34.9942 | -28.1241
17.9587 | -34.9947 | -28.1316
17.9648 | -34.9952 | -28.1392
17.9709 | -34.9956 | -28.1467
17.977 | -34.9962 | -28.1539
17.9832 | -34.9967 | -28.1616
17.9893 | -34.9972 | -28.1689
17.9954 | -34.9976 | -28.1766
18.0015 | -34.9981 | -28.1842
18.0076 | -34.9986 | -28.1916
18.0137 | -34.9991 | -28.1993
18.0198 | -34.9996 | -28.2068
18.0259 | -35.0001 | -28.2145
18.032 | -35.0006 | -28.2223
18.0381 | -35.001 | -28.2298
18.0442 | -35.0015 | -28.2373
18.0503 | -35.002 | -28.2447
18.0564 | -35.0025 | -28.2522
18.0625 | -35.003 | -28.2597
18.0686 | -35.0035 | -28.2673
18.0747 | -35.004 | -28.2748
18.0808 | -35.0045 | -28.282
18.0869 | -35.005 | -28.2895
18.093 | -35.0055 | -28.297
18.0991 | -35.006 | -28.3044
18.1052 | -35.0065 | -28.312
18.1113 | -35.007 | -28.3192
18.1174 | -35.0075 | -28.3264
18.1236 | -35.008 | -28.334
18.1297 | -35.0084 | -28.3416
18.1358 | -35.0089 | -28.3491
18.1419 | -35.0094 | -28.3567
18.148 | -35.0099 | -28.3641
18.1541 | -35.0104 | -28.3715
18.1602 | -35.0109 | -28.379
18.1663 | -35.0114 | -28.3864
18.1724 | -35.0119 | -28.394
18.1785 | -35.0123 | -28.4013
18.1846 | -35.0128 | -28.4087
18.1907 | -35.0133 | -28.4162
18.1968 | -35.0138 | -28.4237
18.2029 | -35.0143 | -28.4313
18.209 | -35.0148 | -28.4389
18.2151 | -35.0152 | -28.4464
18.2212 | -35.0157 | -28.4538
18.2273 | -35.0162 | -28.4613
18.2334 | -35.0167 | -28.4688
18.2395 | -35.0172 | -28.4763
18.2456 | -35.0177 | -28.4838
18.2517 | -35.0181 | -28.4914
18.2578 | -35.0187 | -28.4989
18.2639 | -35.0192 | -28.5064
18.2701 | -35.0197 | -28.514
18.2762 | -35.0202 | -28.5212
18.2823 | -35.0206 | -28.5289
18.2884 | -35.0211 | -28.5362
18.2945 | -35.0216 | -28.5434
18.3006 | -35.0221 | -28.5509
18.3067 | -35.0226 | -28.5583
18.3128 | -35.0231 | -28.5659
18.3189 | -35.0236 | -28.5734
18.325 | -35.0241 | -28.5806
18.3311 | -35.0246 | -28.5879
18.3372 | -35.0251 | -28.5953
18.3433 | -35.0256 | -28.6028
18.3494 | -35.026 | -28.6102
18.3555 | -35.0265 | -28.6176
18.3616 | -35.027 | -28.6249
18.3677 | -35.0275 | -28.6322
18.3738 | -35.028 | -28.6395
18.3799 | -35.0285 | -28.6468
18.386 | -35.029 | -28.6542
18.3921 | -35.0294 | -28.6616
18.3982 | -35.0299 | -28.6688
18.4043 | -35.0304 | -28.6761
18.4105 | -35.0309 | -28.6837
18.4166 | -35.0313 | -28.6908
18.4227 | -35.0318 | -28.6982
18.4288 | -35.0323 | -28.7058
18.4349 | -35.0328 | -28.7128
18.441 | -35.0333 | -28.72
18.4471 | -35.0337 | -28.7274
18.4532 | -35.0342 | -28.7348
18.4593 | -35.0347 | -28.7423
18.4654 | -35.0352 | -28.7494
18.4715 | -35.0357 | -28.7567
18.4776 | -35.0362 | -28.7642
18.4837 | -35.0367 | -28.7716
18.4898 | -35.0372 | -28.779
18.4959 | -35.0377 | -28.7865
18.502 | -35.0382 | -28.794
18.5081 | -35.0386 | -28.8016
18.5142 | -35.0391 | -28.8092
18.5203 | -35.0396 | -28.8165
18.5264 | -35.0401 | -28.8238
18.5325 | -35.0406 | -28.831
18.5386 | -35.0411 | -28.8383
18.5447 | -35.0416 | -28.8455
18.5508 | -35.0421 | -28.8529
18.557 | -35.0426 | -28.8604
18.5631 | -35.0431 | -28.8678
18.5692 | -35.0435 | -28.8751
18.5753 | -35.044 | -28.8822
18.5814 | -35.0445 | -28.8896
18.5875 | -35.045 | -28.8969
18.5936 | -35.0455 | -28.9042
18.5997 | -35.0459 | -28.9117
18.6058 | -35.0464 | -28.919
18.6119 | -35.0469 | -28.9264
18.618 | -35.0473 | -28.9338
18.6241 | -35.0479 | -28.9411
18.6302 | -35.0483 | -28.9483
18.6363 | -35.0488 | -28.9556
18.6424 | -35.0493 | -28.963
18.6485 | -35.0498 | -28.9703
18.6546 | -35.0503 | -28.9777
18.6607 | -35.0507 | -28.985
18.6668 | -35.0512 | -28.9923
18.6729 | -35.0517 | -28.9997
18.679 | -35.0522 | -29.0067
18.6851 | -35.0527 | -29.0139
18.6912 | -35.0532 | -29.0211
18.6974 | -35.0537 | -29.0283
18.7035 | -35.0541 | -29.0357
18.7096 | -35.0546 | -29.0433
18.7157 | -35.0551 | -29.0506
18.7218 | -35.0556 | -29.0578
18.7279 | -35.0561 | -29.0652
18.734 | -35.0566 | -29.0726
18.7401 | -35.0571 | -29.08
18.7462 | -35.0576 | -29.0871
18.7523 | -35.0581 | -29.0945
18.7584 | -35.0586 | -29.102
18.7645 | -35.0591 | -29.1095
18.7706 | -35.0595 | -29.1168
18.7767 | -35.06 | -29.124
18.7828 | -35.0605 | -29.131
18.7889 | -35.061 | -29.1383
18.795 | -35.0615 | -29.1458
18.8011 | -35.0619 | -29.1531
18.8072 | -35.0624 | -29.1606
18.8133 | -35.0629 | -29.1679
18.8194 | -35.0634 | -29.1749
18.8255 | -35.0638 | -29.1823
18.8316 | -35.0643 | -29.1898
18.8377 | -35.0648 | -29.1968
18.8439 | -35.0653 | -29.2041
18.85 | -35.0658 | -29.2115
18.8561 | -35.0663 | -29.2185
18.8622 | -35.0668 | -29.2257
18.8683 | -35.0673 | -29.233
18.8744 | -35.0678 | -29.2403
18.8805 | -35.0683 | -29.2476
18.8866 | -35.0688 | -29.2549
18.8927 | -35.0693 | -29.262
18.8988 | -35.0697 | -29.2693
18.9049 | -35.0702 | -29.2766
18.911 | -35.0707 | -29.2838
18.9171 | -35.0711 | -29.2911
18.9232 | -35.0716 | -29.2981
18.9293 | -35.0721 | -29.3053
18.9354 | -35.0726 | -29.3125
18.9415 | -35.0731 | -29.3196
18.9476 | -35.0736 | -29.3269
18.9537 | -35.074 | -29.3342
18.9598 | -35.0745 | -29.3411
18.9659 | -35.075 | -29.3482
18.972 | -35.0754 | -29.3554
18.9781 | -35.0759 | -29.3628
18.9843 | -35.0763 | -29.37
18.9904 | -35.0769 | -29.3773
18.9965 | -35.0773 | -29.3844
19.0026 | -35.0778 | -29.3916
19.0087 | -35.0784 | -29.3988
19.0148 | -35.0788 | -29.4059
19.0209 | -35.0793 | -29.4132
19.027 | -35.0798 | -29.4203
19.0331 | -35.0803 | -29.4274
19.0392 | -35.0808 | -29.4345
19.0453 | -35.0813 | -29.4416
19.0514 | -35.0818 | -29.4487
19.0575 | -35.0823 | -29.4562
19.0636 | -35.0827 | -29.463
19.0697 | -35.0832 | -29.4702
19.0758 | -35.0837 | -29.4775
19.0819 | -35.0842 | -29.4847
19.088 | -35.0847 | -29.492
19.0941 | -35.0851 | -29.4993
19.1002 | -35.0856 | -29.5062
19.1063 | -35.0861 | -29.5133
19.1124 | -35.0865 | -29.5205
19.1185 | -35.087 | -29.5275
19.1246 | -35.0875 | -29.5346
19.1308 | -35.088 | -29.542
19.1369 | -35.0885 | -29.5492
19.143 | -35.089 | -29.5563
19.1491 | -35.0895 | -29.5634
19.1552 | -35.0899 | -29.5706
19.1613 | -35.0904 | -29.5777
19.1674 | -35.0909 | -29.5851
19.1735 | -35.0914 | -29.5922
19.1796 | -35.0919 | -29.5994
19.1857 | -35.0923 | -29.6065
19.1918 | -35.0928 | -29.6137
19.1979 | -35.0933 | -29.621
19.204 | -35.0938 | -29.6281
19.2101 | -35.0942 | -29.6353
19.2162 | -35.0947 | -29.6424
19.2223 | -35.0952 | -29.6499
19.2284 | -35.0956 | -29.657
19.2345 | -35.0961 | -29.6642
19.2406 | -35.0966 | -29.6715
19.2467 | -35.0971 | -29.6786
19.2528 | -35.0975 | -29.6858
19.2589 | -35.098 | -29.693
19.265 | -35.0985 | -29.7001
19.2712 | -35.099 | -29.7074
19.2773 | -35.0994 | -29.7145
19.2834 | -35.0999 | -29.7217
19.2895 | -35.1004 | -29.729
19.2956 | -35.1009 | -29.7363
19.3017 | -35.1014 | -29.7433
19.3078 | -35.1018 | -29.7504
19.3139 | -35.1023 | -29.7573
19.32 | -35.1028 | -29.7643
19.3261 | -35.1033 | -29.7717
19.3322 | -35.1038 | -29.7787
19.3383 | -35.1043 | -29.7858
19.3444 | -35.1048 | -29.7928
19.3505 | -35.1053 | -29.8002
19.3566 | -35.1058 | -29.8074
19.3627 | -35.1063 | -29.8146
19.3688 | -35.1068 | -29.8217
19.3749 | -35.1072 | -29.8288
19.381 | -35.1077 | -29.8355
19.3871 | -35.1082 | -29.8425
19.3932 | -35.1086 | -29.8497
19.3993 | -35.1091 | -29.8567
19.4054 | -35.1096 | -29.8642
19.4115 | -35.1101 | -29.8712
19.4177 | -35.1105 | -29.8782
19.4238 | -35.111 | -29.8851
19.4299 | -35.1115 | -29.8922
19.436 | -35.112 | -29.8994
19.4421 | -35.1125 | -29.9064
19.4482 | -35.1129 | -29.9136
19.4543 | -35.1134 | -29.9204
19.4604 | -35.1139 | -29.9274
19.4665 | -35.1144 | -29.9347
19.4726 | -35.1149 | -29.9417
19.4787 | -35.1153 | -29.9489
19.4848 | -35.1158 | -29.9561
19.4909 | -35.1162 | -29.9631
19.497 | -35.1167 | -29.97
19.5031 | -35.1172 | -29.977
19.5092 | -35.1176 | -29.9843
19.5153 | -35.1181 | -29.9915
19.5214 | -35.1186 | -29.9984
19.5275 | -35.119 | -30.0055
19.5336 | -35.1195 | -30.0126
19.5397 | -35.12 | -30.0197
19.5458 | -35.1205 | -30.0269
19.5519 | -35.121 | -30.0338
19.5581 | -35.1214 | -30.0406
19.5642 | -35.1219 | -30.0476
19.5703 | -35.1223 | -30.0548
19.5764 | -35.1228 | -30.0621
19.5825 | -35.1233 | -30.0693
19.5886 | -35.1238 | -30.0765
19.5947 | -35.1242 | -30.0837
19.6008 | -35.1246 | -30.0906
19.6069 | -35.1252 | -30.0978
19.613 | -35.1256 | -30.1049
19.6191 | -35.1261 | -30.1119
19.6252 | -35.1265 | -30.1194
19.6313 | -35.127 | -30.1258
19.6374 | -35.1275 | -30.1328
19.6435 | -35.128 | -30.1398
19.6496 | -35.1285 | -30.1468
19.6557 | -35.1289 | -30.154
19.6618 | -35.1294 | -30.1609
19.6679 | -35.1299 | -30.168
19.674 | -35.1304 | -30.1751
19.6801 | -35.1308 | -30.1823
19.6862 | -35.1313 | -30.1894
19.6923 | -35.1318 | -30.1965
19.6984 | -35.1323 | -30.2034
19.7046 | -35.1328 | -30.2104
19.7107 | -35.1332 | -30.2173
19.7168 | -35.1337 | -30.2246
19.7229 | -35.1342 | -30.2316
19.729 | -35.1347 | -30.2385
19.7351 | -35.1351 | -30.2454
19.7412 | -35.1356 | -30.2523
19.7473 | -35.1361 | -30.2592
19.7534 | -35.1365 | -30.2664
19.7595 | -35.137 | -30.2733
19.7656 | -35.1375 | -30.2804
19.7717 | -35.138 | -30.2868
19.7778 | -35.1384 | -30.2938
19.7839 | -35.1389 | -30.3009
19.79 | -35.1394 | -30.308
19.7961 | -35.1398 | -30.315
19.8022 | -35.1403 | -30.322
19.8083 | -35.1407 | -30.3291
19.8144 | -35.1412 | -30.3359
19.8205 | -35.1417 | -30.3429
19.8266 | -35.1422 | -30.3497
19.8327 | -35.1427 | -30.3564
19.8388 | -35.1431 | -30.3631
19.845 | -35.1436 | -30.3702
19.8511 | -35.1441 | -30.3774
19.8572 | -35.1445 | -30.3843
19.8633 | -35.145 | -30.3912
19.8694 | -35.1455 | -30.3981
19.8755 | -35.1459 | -30.4046
19.8816 | -35.1464 | -30.4118
19.8877 | -35.1468 | -30.4188
19.8938 | -35.1473 | -30.4258
19.8999 | -35.1478 | -30.4328
19.906 | -35.1482 | -30.4395
19.9121 | -35.1487 | -30.4463
19.9182 | -35.1492 | -30.4533
19.9243 | -35.1497 | -30.4601
19.9304 | -35.1501 | -30.4672
19.9365 | -35.1506 | -30.4741
19.9426 | -35.151 | -30.4807
19.9487 | -35.1515 | -30.4878
19.9548 | -35.152 | -30.4947
19.9609 | -35.1524 | -30.5014
19.967 | -35.1529 | -30.5082
19.9731 | -35.1533 | -30.5153
19.9792 | -35.1538 | -30.5223
19.9853 | -35.1543 | -30.5294
19.9915 | -35.1548 | -30.5363
19.9976 | -35.1552 | -30.5433
20.0037 | -35.1557 | -30.5504
20.0098 | -35.1562 | -30.5572
20.0159 | -35.1567 | -30.5637
20.022 | -35.1571 | -30.5707
20.0281 | -35.1576 | -30.5776
20.0342 | -35.158 | -30.5844
20.0403 | -35.1585 | -30.5913
20.0464 | -35.159 | -30.5983
20.0525 | -35.1595 | -30.6046
20.0586 | -35.1599 | -30.6118
20.0647 | -35.1604 | -30.6187
20.0708 | -35.1609 | -30.6257
20.0769 | -35.1614 | -30.6327
20.083 | -35.1619 | -30.6391
20.0891 | -35.1623 | -30.646
20.0952 | -35.1628 | -30.653
20.1013 | -35.1632 | -30.6599
20.1074 | -35.1637 | -30.6668
20.1135 | -35.1641 | -30.6736
20.1196 | -35.1646 | -30.6803
20.1257 | -35.165 | -30.6872
20.1319 | -35.1655 | -30.6943
20.138 | -35.166 | -30.7011
20.1441 | -35.1665 | -30.7079
20.1502 | -35.1669 | -30.7147
20.1563 | -35.1674 | -30.7213
20.1624 | -35.1679 | -30.7284
20.1685 | -35.1684 | -30.7354
20.1746 | -35.1689 | -30.7423
20.1807 | -35.1694 | -30.7492
20.1868 | -35.1699 | -30.7564
20.1929 | -35.1704 | -30.7629
20.199 | -35.1708 | -30.7699
20.2051 | -35.1713 | -30.7768
20.2112 | -35.1717 | -30.7836
20.2173 | -35.1722 | -30.7901
20.2234 | -35.1727 | -30.7971
20.2295 | -35.1731 | -30.8038
20.2356 | -35.1736 | -30.8106
20.2417 | -35.1741 | -30.8173
20.2478 | -35.1745 | -30.8242
20.2539 | -35.175 | -30.8312
20.26 | -35.1755 | -30.8378
20.2661 | -35.176 | -30.8446
20.2723 | -35.1765 | -30.8514
20.2784 | -35.1769 | -30.8584
20.2845 | -35.1774 | -30.865
20.2906 | -35.1778 | -30.8719
20.2967 | -35.1783 | -30.8786
20.3028 | -35.1787 | -30.8855
20.3089 | -35.1792 | -30.8925
20.315 | -35.1797 | -30.8992
20.3211 | -35.1801 | -30.9061
20.3272 | -35.1806 | -30.9129
20.3333 | -35.1811 | -30.9194
20.3394 | -35.1815 | -30.9259
20.3455 | -35.182 | -30.9328
20.3516 | -35.1825 | -30.9395
20.3577 | -35.1829 | -30.9464
20.3638 | -35.1834 | -30.9531
20.3699 | -35.1838 | -30.9598
20.376 | -35.1843 | -30.9665
20.3821 | -35.1848 | -30.9734
20.3882 | -35.1853 | -30.9802
20.3943 | -35.1857 | -30.987
20.4004 | -35.1861 | -30.9936
20.4065 | -35.1866 | -31.0002
20.4126 | -35.1871 | -31.0069
20.4188 | -35.1875 | -31.014
20.4249 | -35.188 | -31.0206
20.431 | -35.1884 | -31.0272
20.4371 | -35.1889 | -31.0336
20.4432 | -35.1893 | -31.0402
20.4493 | -35.1898 | -31.0468
20.4554 | -35.1903 | -31.0534
20.4615 | -35.1907 | -31.0602
20.4676 | -35.1911 | -31.067
20.4737 | -35.1916 | -31.0732
20.4798 | -35.192 | -31.0802
20.4859 | -35.1925 | -31.0868
20.492 | -35.193 | -31.0936
20.4981 | -35.1935 | -31.1003
20.5042 | -35.194 | -31.107
20.5103 | -35.1945 | -31.1137
20.5164 | -35.1949 | -31.1203
20.5225 | -35.1954 | -31.127
20.5286 | -35.1958 | -31.1336
20.5347 | -35.1962 | -31.1404
20.5408 | -35.1967 | -31.1466
20.5469 | -35.1971 | -31.1535
20.553 | -35.1976 | -31.1601
20.5592 | -35.1981 | -31.1667
20.5653 | -35.1985 | -31.1736
20.5714 | -35.199 | -31.1805
20.5775 | -35.1995 | -31.1873
20.5836 | -35.1999 | -31.1943
20.5897 | -35.2003 | -31.2014
20.5958 | -35.2008 | -31.2079
20.6019 | -35.2013 | -31.2144
20.608 | -35.2017 | -31.2209
20.6141 | -35.2022 | -31.2273
20.6202 | -35.2026 | -31.234
20.6263 | -35.2031 | -31.2406
20.6324 | -35.2036 | -31.2469
20.6385 | -35.204 | -31.2538
20.6446 | -35.2045 | -31.2603
20.6507 | -35.2049 | -31.2671
20.6568 | -35.2053 | -31.274
20.6629 | -35.2058 | -31.2805
20.669 | -35.2063 | -31.2872
20.6751 | -35.2067 | -31.2937
20.6812 | -35.2072 | -31.2997
20.6873 | -35.2076 | -31.3063
20.6934 | -35.2081 | -31.3127
20.6995 | -35.2085 | -31.319
20.7057 | -35.2089 | -31.3259
20.7118 | -35.2094 | -31.3324
20.7179 | -35.2099 | -31.3387
20.724 | -35.2104 | -31.3453
20.7301 | -35.2108 | -31.3516
20.7362 | -35.2113 | -31.3584
20.7423 | -35.2118 | -31.365
20.7484 | -35.2123 | -31.3714
20.7545 | -35.2127 | -31.3779
20.7606 | -35.2132 | -31.3847
20.7667 | -35.2136 | -31.3913
20.7728 | -35.2141 | -31.3977
20.7789 | -35.2146 | -31.4041
20.785 | -35.2151 | -31.4107
20.7911 | -35.2155 | -31.4167
20.7972 | -35.216 | -31.4232
20.8033 | -35.2164 | -31.4296
20.8094 | -35.2169 | -31.4363
20.8155 | -35.2174 | -31.4425
20.8216 | -35.2178 | -31.449
20.8277 | -35.2183 | -31.4549
20.8338 | -35.2188 | -31.4614
20.8399 | -35.2192 | -31.468
20.8461 | -35.2197 | -31.4747
20.8522 | -35.2202 | -31.4817
20.8583 | -35.2207 | -31.4881
20.8644 | -35.2211 | -31.4944
20.8705 | -35.2216 | -31.5011
20.8766 | -35.2221 | -31.5078
20.8827 | -35.2225 | -31.5144
20.8888 | -35.2229 | -31.5212
20.8949 | -35.2234 | -31.5274
20.901 | -35.2239 | -31.5343
20.9071 | -35.2244 | -31.541
20.9132 | -35.2248 | -31.5474
20.9193 | -35.2252 | -31.554
20.9254 | -35.2257 | -31.5606
20.9315 | -35.2262 | -31.5669
20.9376 | -35.2267 | -31.5731
20.9437 | -35.2272 | -31.5797
20.9498 | -35.2277 | -31.5867
20.9559 | -35.2281 | -31.5929
20.962 | -35.2285 | -31.5992
20.9681 | -35.229 | -31.6057
20.9742 | -35.2295 | -31.6125
20.9803 | -35.2299 | -31.6192
20.9864 | -35.2303 | -31.6257
20.9926 | -35.2309 | -31.6321
20.9987 | -35.2313 | -31.6385
21.0048 | -35.2318 | -31.6453
21.0109 | -35.2322 | -31.652
21.017 | -35.2326 | -31.6583
21.0231 | -35.233 | -31.6646
21.0292 | -35.2336 | -31.6711
21.0353 | -35.234 | -31.6772
21.0414 | -35.2345 | -31.6838
21.0475 | -35.2349 | -31.6907
21.0536 | -35.2354 | -31.6972
21.0597 | -35.2359 | -31.7034
21.0658 | -35.2364 | -31.7099
21.0719 | -35.2369 | -31.7162
21.078 | -35.2373 | -31.7228
21.0841 | -35.2378 | -31.7292
21.0902 | -35.2382 | -31.7353
21.0963 | -35.2388 | -31.7418
21.1024 | -35.2393 | -31.7484
21.1085 | -35.2398 | -31.7552
21.1146 | -35.2403 | -31.7615
21.1207 | -35.2408 | -31.7683
21.1268 | -35.2413 | -31.7747
21.133 | -35.2417 | -31.7813
21.1391 | -35.2422 | -31.7877
21.1452 | -35.2427 | -31.7937
21.1513 | -35.2431 | -31.8001
21.1574 | -35.2435 | -31.8065
21.1635 | -35.244 | -31.8128
21.1696 | -35.2445 | -31.8191
21.1757 | -35.245 | -31.825
21.1818 | -35.2454 | -31.8316
21.1879 | -35.2459 | -31.8379
21.194 | -35.2464 | -31.8442
21.2001 | -35.2468 | -31.8508
21.2062 | -35.2472 | -31.8572
21.2123 | -35.2477 | -31.8634
21.2184 | -35.2482 | -31.8698
21.2245 | -35.2485 | -31.8764
21.2306 | -35.249 | -31.8829
21.2367 | -35.2495 | -31.8896
21.2428 | -35.25 | -31.8962
21.2489 | -35.2504 | -31.9027
21.255 | -35.2508 | -31.9092
21.2611 | -35.2513 | -31.916
21.2672 | -35.2517 | -31.9222
21.2733 | -35.2521 | -31.9286
21.2795 | -35.2526 | -31.9346
21.2856 | -35.2531 | -31.9408
21.2917 | -35.2536 | -31.9474
21.2978 | -35.254 | -31.9531
21.3039 | -35.2545 | -31.9597
21.31 | -35.255 | -31.9659
21.3161 | -35.2554 | -31.9721
21.3222 | -35.2559 | -31.9786
21.3283 | -35.2563 | -31.9855
21.3344 | -35.2567 | -31.992
21.3405 | -35.2572 | -31.9981
21.3466 | -35.2578 | -32.0046
21.3527 | -35.2582 | -32.0114
21.3588 | -35.2586 | -32.0176
21.3649 | -35.2591 | -32.0238
21.371 | -35.2596 | -32.0304
21.3771 | -35.26 | -32.0367
21.3832 | -35.2605 | -32.0427
21.3893 | -35.261 | -32.0486
21.3954 | -35.2614 | -32.0553
21.4015 | -35.2619 | -32.0617
21.4076 | -35.2624 | -32.0682
21.4137 | -35.2628 | -32.0742
21.4199 | -35.2633 | -32.0805
21.426 | -35.2637 | -32.0865
21.4321 | -35.2641 | -32.0931
21.4382 | -35.2646 | -32.0995
21.4443 | -35.265 | -32.1056
21.4504 | -35.2655 | -32.1122
21.4565 | -35.266 | -32.1184
21.4626 | -35.2665 | -32.1244
21.4687 | -35.2669 | -32.1304
21.4748 | -35.2674 | -32.1364
21.4809 | -35.2678 | -32.1427
21.487 | -35.2683 | -32.1489
21.4931 | -35.2688 | -32.1551
21.4992 | -35.2692 | -32.1615
21.5053 | -35.2697 | -32.1679
21.5114 | -35.2702 | -32.1743
21.5175 | -35.2706 | -32.1809
21.5236 | -35.271 | -32.1875
21.5297 | -35.2715 | -32.194
21.5358 | -35.272 | -32.2001
21.5419 | -35.2724 | -32.2065
21.548 | -35.2728 | -32.213
21.5541 | -35.2733 | -32.2195
21.5602 | -35.2737 | -32.226
21.5664 | -35.2742 | -32.2322
21.5725 | -35.2746 | -32.2382
21.5786 | -35.275 | -32.2442
21.5847 | -35.2754 | -32.2505
21.5908 | -35.276 | -32.2568
21.5969 | -35.2764 | -32.2626
21.603 | -35.2769 | -32.2688
21.6091 | -35.2773 | -32.2752
21.6152 | -35.2777 | -32.2814
21.6213 | -35.2782 | -32.2872
21.6274 | -35.2787 | -32.2936
21.6335 | -35.279 | -32.2998
21.6396 | -35.2794 | -32.3061
21.6457 | -35.2799 | -32.3122
21.6518 | -35.2804 | -32.3184
21.6579 | -35.2808 | -32.3243
21.664 | -35.2813 | -32.3309
21.6701 | -35.2817 | -32.3365
21.6762 | -35.2821 | -32.3426
21.6823 | -35.2825 | -32.349
21.6884 | -35.283 | -32.3553
21.6945 | -35.2835 | -32.3616
21.7006 | -35.2839 | -32.3678
21.7068 | -35.2844 | -32.3739
21.7129 | -35.2848 | -32.3798
21.719 | -35.2853 | -32.3862
21.7251 | -35.2858 | -32.3923
21.7312 | -35.2862 | -32.3985
21.7373 | -35.2866 | -32.4048
21.7434 | -35.2871 | -32.411
21.7495 | -35.2876 | -32.4175
21.7556 | -35.288 | -32.4236
21.7617 | -35.2885 | -32.4303
21.7678 | -35.289 | -32.4369
21.7739 | -35.2895 | -32.4427
21.78 | -35.2899 | -32.4493
21.7861 | -35.2904 | -32.4554
21.7922 | -35.2908 | -32.4614
21.7983 | -35.2913 | -32.4678
21.8044 | -35.2918 | -32.4743
21.8105 | -35.2922 | -32.4808
21.8166 | -35.2927 | -32.4866
21.8227 | -35.2931 | -32.4927
21.8288 | -35.2936 | -32.499
21.8349 | -35.2941 | -32.5049
21.841 | -35.2945 | -32.5113
21.8471 | -35.2949 | -32.5174
21.8533 | -35.2954 | -32.5238
21.8594 | -35.2958 | -32.5302
21.8655 | -35.2963 | -32.5362
21.8716 | -35.2968 | -32.5424
21.8777 | -35.2972 | -32.548
21.8838 | -35.2976 | -32.554
21.8899 | -35.2981 | -32.5598
21.896 | -35.2985 | -32.5658
21.9021 | -35.299 | -32.5721
21.9082 | -35.2995 | -32.5786
21.9143 | -35.2999 | -32.5844
21.9204 | -35.3004 | -32.5905
21.9265 | -35.3009 | -32.5963
21.9326 | -35.3014 | -32.6028
21.9387 | -35.3018 | -32.609
21.9448 | -35.3023 | -32.6152
21.9509 | -35.3027 | -32.6214
21.957 | -35.3032 | -32.6276
21.9631 | -35.3037 | -32.6339
21.9692 | -35.3041 | -32.6395
21.9753 | -35.3045 | -32.6455
21.9814 | -35.305 | -32.6516
21.9875 | -35.3054 | -32.6569
21.9937 | -35.3058 | -32.6629
21.9998 | -35.3063 | -32.6689
22.0059 | -35.3067 | -32.6751
22.012 | -35.3072 | -32.6809
22.0181 | -35.3076 | -32.6869
22.0242 | -35.3081 | -32.6933
22.0303 | -35.3086 | -32.6993
22.0364 | -35.309 | -32.7056
22.0425 | -35.3095 | -32.7114
22.0486 | -35.3101 | -32.7171
22.0547 | -35.3105 | -32.7229
22.0608 | -35.311 | -32.7291
22.0669 | -35.3114 | -32.7356
22.073 | -35.3119 | -32.742
22.0791 | -35.3124 | -32.7478
22.0852 | -35.3129 | -32.7541
22.0913 | -35.3134 | -32.7599
22.0974 | -35.3138 | -32.7663
22.1035 | -35.3142 | -32.7727
22.1096 | -35.3146 | -32.779
22.1157 | -35.3151 | -32.7851
22.1218 | -35.3155 | -32.7912
22.1279 | -35.316 | -32.7973
22.134 | -35.3165 | -32.8035
22.1402 | -35.317 | -32.8096
22.1463 | -35.3174 | -32.8151
22.1524 | -35.3178 | -32.8211
22.1585 | -35.3183 | -32.8271
22.1646 | -35.3187 | -32.833
22.1707 | -35.3192 | -32.8391
22.1768 | -35.3196 | -32.8451
22.1829 | -35.32 | -32.8511
22.189 | -35.3205 | -32.8568
22.1951 | -35.321 | -32.8633
22.2012 | -35.3214 | -32.8693
22.2073 | -35.3219 | -32.8751
22.2134 | -35.3223 | -32.8812
22.2195 | -35.3227 | -32.8871
22.2256 | -35.3231 | -32.8933
22.2317 | -35.3236 | -32.8991
22.2378 | -35.324 | -32.9048
22.2439 | -35.3245 | -32.9111
22.25 | -35.3249 | -32.9172
22.2561 | -35.3253 | -32.9231
22.2622 | -35.3257 | -32.9292
22.2683 | -35.3262 | -32.9347
22.2744 | -35.3266 | -32.9404
22.2806 | -35.3271 | -32.947
22.2867 | -35.3276 | -32.953
22.2928 | -35.328 | -32.9591
22.2989 | -35.3285 | -32.9651
22.305 | -35.3289 | -32.9709
22.3111 | -35.3294 | -32.9773
22.3172 | -35.3298 | -32.9831
22.3233 | -35.3303 | -32.9889
22.3294 | -35.3308 | -32.995
22.3355 | -35.3312 | -33.001
22.3416 | -35.3317 | -33.0069
22.3477 | -35.3321 | -33.0128
22.3538 | -35.3325 | -33.0193
22.3599 | -35.333 | -33.0252
22.366 | -35.3335 | -33.0314
22.3721 | -35.3339 | -33.0372
22.3782 | -35.3344 | -33.043
22.3843 | -35.3348 | -33.0492
22.3904 | -35.3352 | -33.0547
22.3965 | -35.3357 | -33.0612
22.4026 | -35.3361 | -33.0673
22.4087 | -35.3366 | -33.073
22.4148 | -35.337 | -33.0791
22.4209 | -35.3375 | -33.0848
22.4271 | -35.338 | -33.0909
22.4332 | -35.3384 | -33.0972
22.4393 | -35.3388 | -33.1033
22.4454 | -35.3392 | -33.1086
22.4515 | -35.3397 | -33.1146
22.4576 | -35.3402 | -33.1207
22.4637 | -35.3406 | -33.1267
22.4698 | -35.3411 | -33.1328
22.4759 | -35.3415 | -33.1384
22.482 | -35.3419 | -33.1446
22.4881 | -35.3423 | -33.1503
22.4942 | -35.3426 | -33.1564
22.5003 | -35.3431 | -33.1617
22.5064 | -35.3435 | -33.1682
22.5125 | -35.3438 | -33.174
22.5186 | -35.3443 | -33.1802
22.5247 | -35.3448 | -33.1863
22.5308 | -35.3452 | -33.1916
22.5369 | -35.3456 | -33.1977
22.543 | -35.3461 | -33.2034
22.5491 | -35.3465 | -33.2086
22.5552 | -35.3469 | -33.2146
22.5613 | -35.3474 | -33.2203
22.5675 | -35.3479 | -33.2262
22.5736 | -35.3482 | -33.2324
22.5797 | -35.3487 | -33.2377
22.5858 | -35.3491 | -33.2437
22.5919 | -35.3495 | -33.2496
22.598 | -35.35 | -33.2556
22.6041 | -35.3504 | -33.2615
22.6102 | -35.3509 | -33.2673
22.6163 | -35.3513 | -33.2735
22.6224 | -35.3517 | -33.2799
22.6285 | -35.3521 | -33.2856
22.6346 | -35.3526 | -33.2911
22.6407 | -35.353 | -33.297
22.6468 | -35.3534 | -33.3028
22.6529 | -35.3539 | -33.3082
22.659 | -35.3544 | -33.3143
22.6651 | -35.3549 | -33.3203
22.6712 | -35.3554 | -33.3264
22.6773 | -35.3558 | -33.3323
22.6834 | -35.3563 | -33.3386
22.6895 | -35.3568 | -33.3441
22.6956 | -35.3572 | -33.3502
22.7017 | -35.3577 | -33.3562
22.7079 | -35.3581 | -33.3618
22.714 | -35.3585 | -33.3679
22.7201 | -35.3589 | -33.3737
22.7262 | -35.3595 | -33.38
22.7323 | -35.3599 | -33.3859
22.7384 | -35.3604 | -33.3921
22.7445 | -35.3609 | -33.398
22.7506 | -35.3613 | -33.404
22.7567 | -35.3618 | -33.41
22.7628 | -35.3622 | -33.416
22.7689 | -35.3626 | -33.4224
22.775 | -35.3631 | -33.4284
22.7811 | -35.3635 | -33.4346
22.7872 | -35.364 | -33.4406
22.7933 | -35.3644 | -33.4466
22.7994 | -35.3649 | -33.4522
22.8055 | -35.3655 | -33.4579
22.8116 | -35.3659 | -33.4638
22.8177 | -35.3664 | -33.4701
22.8238 | -35.3669 | -33.4757
22.8299 | -35.3674 | -33.4813
22.836 | -35.3679 | -33.4871
22.8421 | -35.3683 | -33.4927
22.8482 | -35.3687 | -33.4988
22.8544 | -35.3692 | -33.5045
22.8605 | -35.3697 | -33.5102
22.8666 | -35.37 | -33.5155
22.8727 | -35.3705 | -33.5215
22.8788 | -35.3709 | -33.5275
22.8849 | -35.3713 | -33.533
22.891 | -35.3717 | -33.5392
22.8971 | -35.3722 | -33.5452
22.9032 | -35.3726 | -33.5516
22.9093 | -35.3731 | -33.5575
22.9154 | -35.3736 | -33.5628
22.9215 | -35.374 | -33.5684
22.9276 | -35.3744 | -33.574
22.9337 | -35.3749 | -33.5799
22.9398 | -35.3753 | -33.5863
22.9459 | -35.3757 | -33.5919
22.952 | -35.3762 | -33.5978
22.9581 | -35.3767 | -33.6036
22.9642 | -35.3772 | -33.6096
22.9703 | -35.3777 | -33.615
22.9764 | -35.3781 | -33.621
22.9825 | -35.3786 | -33.6268
22.9886 | -35.3791 | -33.6328
22.9948 | -35.3795 | -33.6387
23.0009 | -35.38 | -33.6444
23.007 | -35.3804 | -33.6502
23.0131 | -35.3808 | -33.6562
23.0192 | -35.3812 | -33.6617
23.0253 | -35.3816 | -33.6681
23.0314 | -35.3821 | -33.6739
23.0375 | -35.3826 | -33.6796
23.0436 | -35.3831 | -33.6852
23.0497 | -35.3835 | -33.6911
23.0558 | -35.3839 | -33.6971
23.0619 | -35.3844 | -33.703
23.068 | -35.3848 | -33.7092
23.0741 | -35.3852 | -33.7143
23.0802 | -35.3857 | -33.72
23.0863 | -35.3862 | -33.7259
23.0924 | -35.3866 | -33.7318
23.0985 | -35.3871 | -33.7379
23.1046 | -35.3875 | -33.7441
23.1107 | -35.3881 | -33.7495
23.1168 | -35.3885 | -33.7556
23.1229 | -35.389 | -33.7613
23.129 | -35.3895 | -33.7673
23.1351 | -35.3899 | -33.7732
23.1413 | -35.3903 | -33.7786
23.1474 | -35.3908 | -33.7841
23.1535 | -35.3912 | -33.7898
23.1596 | -35.3917 | -33.7959
23.1657 | -35.3921 | -33.8017
23.1718 | -35.3925 | -33.807
23.1779 | -35.3929 | -33.8126
23.184 | -35.3933 | -33.8184
23.1901 | -35.3938 | -33.8246
23.1962 | -35.3942 | -33.8304
23.2023 | -35.3946 | -33.8367
23.2084 | -35.395 | -33.8426
23.2145 | -35.3955 | -33.8482
23.2206 | -35.3958 | -33.8541
23.2267 | -35.3963 | -33.8602
23.2328 | -35.3967 | -33.866
23.2389 | -35.3972 | -33.8717
23.245 | -35.3976 | -33.8778
23.2511 | -35.3981 | -33.8836
23.2572 | -35.3985 | -33.8898
23.2633 | -35.3989 | -33.8959
23.2694 | -35.3994 | -33.9018
23.2755 | -35.3998 | -33.9078
23.2817 | -35.4002 | -33.9135
23.2878 | -35.4007 | -33.9193
23.2939 | -35.4011 | -33.925
23.3 | -35.4016 | -33.9303
23.3061 | -35.402 | -33.9359
23.3122 | -35.4024 | -33.9419
23.3183 | -35.4029 | -33.9477
23.3244 | -35.4033 | -33.9539
23.3305 | -35.4037 | -33.9598
23.3366 | -35.4042 | -33.9656
23.3427 | -35.4046 | -33.9714
23.3488 | -35.405 | -33.9772
23.3549 | -35.4055 | -33.9831
23.361 | -35.4059 | -33.9885
23.3671 | -35.4063 | -33.9945
23.3732 | -35.4067 | -34.0004
23.3793 | -35.4071 | -34.0063
23.3854 | -35.4076 | -34.0124
23.3915 | -35.4081 | -34.0183
23.3976 | -35.4086 | -34.024
23.4037 | -35.409 | -34.0295
23.4098 | -35.4094 | -34.0351
23.4159 | -35.4098 | -34.0412
23.422 | -35.4103 | -34.0471
23.4282 | -35.4107 | -34.0533
23.4343 | -35.4112 | -34.059
23.4404 | -35.4116 | -34.0652
23.4465 | -35.412 | -34.0707
23.4526 | -35.4124 | -34.0763
23.4587 | -35.4129 | -34.0817
23.4648 | -35.4132 | -34.0873
23.4709 | -35.4137 | -34.0927
23.477 | -35.4141 | -34.0986
23.4831 | -35.4146 | -34.1047
23.4892 | -35.415 | -34.1103
23.4953 | -35.4155 | -34.116
23.5014 | -35.4158 | -34.122
23.5075 | -35.4163 | -34.1282
23.5136 | -35.4168 | -34.1338
23.5197 | -35.4172 | -34.1393
23.5258 | -35.4177 | -34.1453
23.5319 | -35.4181 | -34.1511
23.538 | -35.4185 | -34.1567
23.5441 | -35.4189 | -34.1625
23.5502 | -35.4193 | -34.168
23.5563 | -35.4197 | -34.1739
23.5624 | -35.4201 | -34.1796
23.5686 | -35.4206 | -34.1853
23.5747 | -35.4211 | -34.1911
23.5808 | -35.4215 | -34.1967
23.5869 | -35.4219 | -34.2027
23.593 | -35.4223 | -34.2085
23.5991 | -35.4228 | -34.2144
23.6052 | -35.4232 | -34.2202
23.6113 | -35.4237 | -34.2261
23.6174 | -35.4242 | -34.2326
23.6235 | -35.4245 | -34.2386
23.6296 | -35.425 | -34.2443
23.6357 | -35.4255 | -34.2501
23.6418 | -35.4259 | -34.256
23.6479 | -35.4263 | -34.2619
23.654 | -35.4267 | -34.2679
23.6601 | -35.4271 | -34.2737
23.6662 | -35.4276 | -34.2796
23.6723 | -35.4281 | -34.2854
23.6784 | -35.4285 | -34.2912
23.6845 | -35.4289 | -34.297
23.6906 | -35.4294 | -34.3028
23.6967 | -35.4298 | -34.3084
23.7028 | -35.4303 | -34.3145
23.7089 | -35.4307 | -34.3207
23.7151 | -35.431 | -34.3263
23.7212 | -35.4315 | -34.3325
23.7273 | -35.4319 | -34.3379
23.7334 | -35.4324 | -34.3439
23.7395 | -35.4328 | -34.3494
23.7456 | -35.4332 | -34.3549
23.7517 | -35.4336 | -34.3603
23.7578 | -35.434 | -34.366
23.7639 | -35.4344 | -34.3724
23.77 | -35.4348 | -34.3783
23.7761 | -35.4353 | -34.3839
23.7822 | -35.4357 | -34.3891
23.7883 | -35.4361 | -34.3956
23.7944 | -35.4366 | -34.4013
23.8005 | -35.437 | -34.4071
23.8066 | -35.4374 | -34.4132
23.8127 | -35.4379 | -34.419
23.8188 | -35.4383 | -34.4252
23.8249 | -35.4387 | -34.4309
23.831 | -35.4391 | -34.4363
23.8371 | -35.4394 | -34.4426
23.8432 | -35.4399 | -34.449
23.8493 | -35.4402 | -34.4546
23.8555 | -35.4406 | -34.4603
23.8616 | -35.441 | -34.4664
23.8677 | -35.4414 | -34.4721
23.8738 | -35.4419 | -34.478
23.8799 | -35.4423 | -34.4839
23.886 | -35.4427 | -34.4894
23.8921 | -35.4431 | -34.4951
23.8982 | -35.4435 | -34.5009
23.9043 | -35.4439 | -34.5067
23.9104 | -35.4443 | -34.5127
23.9165 | -35.4448 | -34.5189
23.9226 | -35.4451 | -34.5241
23.9287 | -35.4456 | -34.5301
23.9348 | -35.446 | -34.5362
23.9409 | -35.4464 | -34.5421
23.947 | -35.4468 | -34.5482
23.9531 | -35.4473 | -34.5539
23.9592 | -35.4477 | -34.56
23.9653 | -35.4481 | -34.5658
23.9714 | -35.4485 | -34.5715
23.9775 | -35.4489 | -34.5773
23.9836 | -35.4493 | -34.5831
23.9897 | -35.4497 | -34.5888
23.9958 | -35.4501 | -34.5946
24.002 | -35.4506 | -34.6002
24.0081 | -35.4509 | -34.6062
24.0142 | -35.4514 | -34.6121
24.0203 | -35.4518 | -34.6177
24.0264 | -35.4522 | -34.6236
24.0325 | -35.4526 | -34.6294
24.0386 | -35.453 | -34.6351
24.0447 | -35.4534 | -34.6411
24.0508 | -35.4538 | -34.6469
24.0569 | -35.4543 | -34.6528
24.063 | -35.4547 | -34.6582
24.0691 | -35.4551 | -34.6642
24.0752 | -35.4556 | -34.6702
24.0813 | -35.456 | -34.6758
24.0874 | -35.4565 | -34.6812
24.0935 | -35.4569 | -34.6868
24.0996 | -35.4573 | -34.692
24.1057 | -35.4577 | -34.698
24.1118 | -35.4581 | -34.7041
24.1179 | -35.4586 | -34.71
24.124 | -35.4589 | -34.7156
24.1301 | -35.4594 | -34.7211
24.1362 | -35.4598 | -34.7264
24.1424 | -35.4603 | -34.7322
24.1485 | -35.4607 | -34.7386
24.1546 | -35.4611 | -34.7448
24.1607 | -35.4615 | -34.7504
24.1668 | -35.462 | -34.7566
24.1729 | -35.4623 | -34.7628
24.179 | -35.4628 | -34.7689
24.1851 | -35.4632 | -34.7748
24.1912 | -35.4636 | -34.7807
24.1973 | -35.464 | -34.7869
24.2034 | -35.4645 | -34.7923
24.2095 | -35.4649 | -34.7978
24.2156 | -35.4652 | -34.8035
24.2217 | -35.4656 | -34.8094
24.2278 | -35.466 | -34.815
24.2339 | -35.4665 | -34.8213
24.24 | -35.4669 | -34.8267
24.2461 | -35.4673 | -34.8318
24.2522 | -35.4678 | -34.8377
24.2583 | -35.4681 | -34.8432
24.2644 | -35.4686 | -34.8485
24.2705 | -35.469 | -34.854
24.2766 | -35.4694 | -34.8595
24.2827 | -35.4698 | -34.8656
24.2889 | -35.4702 | -34.8718
24.295 | -35.4706 | -34.8775
24.3011 | -35.471 | -34.883
24.3072 | -35.4715 | -34.8889
24.3133 | -35.4719 | -34.8944
24.3194 | -35.4723 | -34.9002
24.3255 | -35.4727 | -34.9064
24.3316 | -35.4731 | -34.9126
24.3377 | -35.4735 | -34.9187
24.3438 | -35.4739 | -34.9243
24.3499 | -35.4743 | -34.9298
24.356 | -35.4747 | -34.9354
24.3621 | -35.4751 | -34.9412
24.3682 | -35.4755 | -34.9472
24.3743 | -35.476 | -34.9527
24.3804 | -35.4763 | -34.9587
24.3865 | -35.4767 | -34.9642
24.3926 | -35.4772 | -34.9702
24.3987 | -35.4776 | -34.9764
24.4048 | -35.4779 | -34.982
24.4109 | -35.4782 | -34.9874
24.417 | -35.4786 | -34.9936
24.4231 | -35.4789 | -34.9992
24.4293 | -35.4793 | -35.0048
24.4354 | -35.4798 | -35.011
24.4415 | -35.4802 | -35.0171
24.4476 | -35.4804 | -35.0238
24.4537 | -35.4808 | -35.0302
24.4598 | -35.4813 | -35.0361
24.4659 | -35.4816 | -35.042
24.472 | -35.482 | -35.0478
24.4781 | -35.4823 | -35.0532
24.4842 | -35.4827 | -35.0594
24.4903 | -35.4831 | -35.0652
24.4964 | -35.4835 | -35.0712
24.5025 | -35.4839 | -35.0771
24.5086 | -35.4843 | -35.0829
24.5147 | -35.4847 | -35.0886
24.5208 | -35.4852 | -35.0949
24.5269 | -35.4856 | -35.1002
24.533 | -35.486 | -35.1059
24.5391 | -35.4864 | -35.1112
24.5452 | -35.4869 | -35.117
24.5513 | -35.4875 | -35.1233
24.5574 | -35.4878 | -35.1289
24.5635 | -35.4883 | -35.1345
24.5696 | -35.4887 | -35.1406
24.5758 | -35.4891 | -35.1465
24.5819 | -35.4895 | -35.152
24.588 | -35.4899 | -35.1583
24.5941 | -35.4903 | -35.1644
24.6002 | -35.4908 | -35.1709
24.6063 | -35.4912 | -35.1764
24.6124 | -35.4916 | -35.1813
24.6185 | -35.4919 | -35.1869
24.6246 | -35.4923 | -35.1922
24.6307 | -35.4927 | -35.1974
24.6368 | -35.4931 | -35.204
24.6429 | -35.4935 | -35.2101
24.649 | -35.494 | -35.2161
24.6551 | -35.4944 | -35.2224
24.6612 | -35.4948 | -35.2285
24.6673 | -35.4952 | -35.2343
24.6734 | -35.4957 | -35.2405
24.6795 | -35.4961 | -35.2463
24.6856 | -35.4966 | -35.2525
24.6917 | -35.497 | -35.2581
24.6978 | -35.4975 | -35.2637
24.7039 | -35.4979 | -35.2699
24.71 | -35.4982 | -35.2759
24.7162 | -35.4987 | -35.2818
24.7223 | -35.4991 | -35.288
24.7284 | -35.4995 | -35.2936
24.7345 | -35.4999 | -35.2991
24.7406 | -35.5004 | -35.3052
24.7467 | -35.5008 | -35.311
24.7528 | -35.5012 | -35.3168
24.7589 | -35.5016 | -35.3226
24.765 | -35.5021 | -35.3289
24.7711 | -35.5025 | -35.3347
24.7772 | -35.5028 | -35.3405
24.7833 | -35.5032 | -35.3465
24.7894 | -35.5035 | -35.3525
24.7955 | -35.5039 | -35.3581
24.8016 | -35.5042 | -35.3638
24.8077 | -35.5047 | -35.3694
24.8138 | -35.5051 | -35.3752
24.8199 | -35.5055 | -35.3811
24.826 | -35.506 | -35.3868
24.8321 | -35.5064 | -35.3924
24.8382 | -35.5067 | -35.3981
24.8443 | -35.5071 | -35.4038
24.8504 | -35.5076 | -35.4095
24.8565 | -35.508 | -35.4154
24.8627 | -35.5084 | -35.4213
24.8688 | -35.5088 | -35.4268
24.8749 | -35.5092 | -35.4326
24.881 | -35.5096 | -35.4387
24.8871 | -35.51 | -35.4444
24.8932 | -35.5104 | -35.4501
24.8993 | -35.5107 | -35.4559
24.9054 | -35.5111 | -35.4622
24.9115 | -35.5116 | -35.4682
24.9176 | -35.512 | -35.474
24.9237 | -35.5124 | -35.4799
24.9298 | -35.5128 | -35.4858
24.9359 | -35.5131 | -35.4916
24.942 | -35.5136 | -35.4976
24.9481 | -35.514 | -35.5034
24.9542 | -35.5145 | -35.5092
24.9603 | -35.5148 | -35.5149
24.9664 | -35.5152 | -35.5209
24.9725 | -35.5155 | -35.527
24.9786 | -35.5159 | -35.5328
24.9847 | -35.5163 | -35.5388
24.9908 | -35.5167 | -35.5445
24.9969 | -35.5171 | -35.5502
25.0031 | -35.5175 | -35.5563
25.0092 | -35.5178 | -35.5621
25.0153 | -35.5182 | -35.5678
25.0214 | -35.5186 | -35.5737
25.0275 | -35.5189 | -35.5795
25.0336 | -35.5193 | -35.5853
25.0397 | -35.5198 | -35.5912
25.0458 | -35.5203 | -35.5967
25.0519 | -35.5207 | -35.6032
25.058 | -35.521 | -35.6094
25.0641 | -35.5214 | -35.6148
25.0702 | -35.5218 | -35.6207
25.0763 | -35.5223 | -35.6267
25.0824 | -35.5226 | -35.6322
25.0885 | -35.523 | -35.6381
25.0946 | -35.5234 | -35.6436
25.1007 | -35.5237 | -35.6495
25.1068 | -35.5241 | -35.6555
25.1129 | -35.5246 | -35.6613
25.119 | -35.525 | -35.6669
25.1251 | -35.5253 | -35.6729
25.1312 | -35.5258 | -35.6787
25.1373 | -35.5262 | -35.6843
25.1435 | -35.5267 | -35.6899
25.1496 | -35.527 | -35.6956
25.1557 | -35.5274 | -35.7013
25.1618 | -35.5277 | -35.7074
25.1679 | -35.5281 | -35.7139
25.174 | -35.5284 | -35.7198
25.1801 | -35.5288 | -35.7257
25.1862 | -35.5293 | -35.7305
25.1923 | -35.5297 | -35.7366
25.1984 | -35.5302 | -35.743
25.2045 | -35.5306 | -35.7478
25.2106 | -35.531 | -35.7532
25.2167 | -35.5314 | -35.7587
25.2228 | -35.5317 | -35.7639
25.2289 | -35.5321 | -35.7698
25.235 | -35.5325 | -35.7757
25.2411 | -35.5329 | -35.7818
25.2472 | -35.5334 | -35.7873
25.2533 | -35.5339 | -35.7932
25.2594 | -35.5342 | -35.7992
25.2655 | -35.5347 | -35.8049
25.2716 | -35.535 | -35.8106
25.2777 | -35.5353 | -35.8162
25.2838 | -35.5358 | -35.8219
25.29 | -35.5362 | -35.8275
25.2961 | -35.5367 | -35.8332
25.3022 | -35.5371 | -35.8391
25.3083 | -35.5374 | -35.8447
25.3144 | -35.5377 | -35.8501
25.3205 | -35.5381 | -35.8553
25.3266 | -35.5385 | -35.8613
25.3327 | -35.5389 | -35.868
25.3388 | -35.5392 | -35.8741
25.3449 | -35.5395 | -35.8795
25.351 | -35.54 | -35.8856
25.3571 | -35.5404 | -35.8915
25.3632 | -35.5409 | -35.8971
25.3693 | -35.5414 | -35.9036
25.3754 | -35.5418 | -35.9096
25.3815 | -35.5422 | -35.9152
25.3876 | -35.5427 | -35.921
25.3937 | -35.543 | -35.9265
25.3998 | -35.5434 | -35.9322
25.4059 | -35.5438 | -35.9378
25.412 | -35.5441 | -35.9438
25.4181 | -35.5445 | -35.9499
25.4242 | -35.5449 | -35.956
25.4304 | -35.5453 | -35.9614
25.4365 | -35.5457 | -35.9676
25.4426 | -35.546 | -35.9732
25.4487 | -35.5464 | -35.9795
25.4548 | -35.5468 | -35.9849
25.4609 | -35.5473 | -35.9903
25.467 | -35.5477 | -35.9959
25.4731 | -35.5481 | -36.0017
25.4792 | -35.5484 | -36.0076
25.4853 | -35.5489 | -36.0134
25.4914 | -35.5493 | -36.0193
25.4975 | -35.5497 | -36.0252
25.5036 | -35.5501 | -36.031
25.5097 | -35.5506 | -36.0371
25.5158 | -35.5509 | -36.0429
25.5219 | -35.5513 | -36.0488
25.528 | -35.5517 | -36.0548
25.5341 | -35.552 | -36.0607
25.5402 | -35.5525 | -36.066
25.5463 | -35.5529 | -36.0721
25.5524 | -35.5533 | -36.078
25.5585 | -35.5537 | -36.0836
25.5646 | -35.5541 | -36.0893
25.5707 | -35.5545 | -36.0954
25.5769 | -35.5548 | -36.102
25.583 | -35.5553 | -36.1078
25.5891 | -35.5557 | -36.1135
25.5952 | -35.556 | -36.1193
25.6013 | -35.5564 | -36.125
25.6074 | -35.5568 | -36.1301
25.6135 | -35.5573 | -36.1362
25.6196 | -35.5577 | -36.1421
25.6257 | -35.5581 | -36.148
25.6318 | -35.5584 | -36.1532
25.6379 | -35.5588 | -36.1584
25.644 | -35.5592 | -36.1634
25.6501 | -35.5595 | -36.1698
25.6562 | -35.5599 | -36.1756
25.6623 | -35.5603 | -36.1817
25.6684 | -35.5607 | -36.1875
25.6745 | -35.5611 | -36.1931
25.6806 | -35.5615 | -36.1988
25.6867 | -35.5618 | -36.2045
25.6928 | -35.5622 | -36.2104
25.6989 | -35.5625 | -36.2159
25.705 | -35.5629 | -36.222
25.7111 | -35.5633 | -36.2277
25.7173 | -35.5638 | -36.234
25.7234 | -35.5641 | -36.2398
25.7295 | -35.5645 | -36.2457
25.7356 | -35.5648 | -36.2512
25.7417 | -35.5652 | -36.2568
25.7478 | -35.5655 | -36.2623
25.7539 | -35.5659 | -36.2685
25.76 | -35.5663 | -36.2744
25.7661 | -35.5667 | -36.2796
25.7722 | -35.5671 | -36.2855
25.7783 | -35.5675 | -36.2912
25.7844 | -35.5678 | -36.297
25.7905 | -35.5682 | -36.3027
25.7966 | -35.5686 | -36.3081
25.8027 | -35.569 | -36.3136
25.8088 | -35.5693 | -36.3194
25.8149 | -35.5697 | -36.3254
25.821 | -35.5701 | -36.3313
25.8271 | -35.5705 | -36.3375
25.8332 | -35.5709 | -36.343
25.8393 | -35.5713 | -36.3489
25.8454 | -35.5717 | -36.3549
25.8515 | -35.5721 | -36.3607
25.8576 | -35.5725 | -36.3664
25.8638 | -35.5729 | -36.3723
25.8699 | -35.5733 | -36.3783
25.876 | -35.5736 | -36.3839
25.8821 | -35.574 | -36.3897
25.8882 | -35.5743 | -36.3955
25.8943 | -35.5747 | -36.4014
25.9004 | -35.575 | -36.4069
25.9065 | -35.5754 | -36.4122
25.9126 | -35.5758 | -36.4178
25.9187 | -35.5761 | -36.4235
25.9248 | -35.5764 | -36.43
25.9309 | -35.5767 | -36.4357
25.937 | -35.5771 | -36.4414
25.9431 | -35.5775 | -36.4471
25.9492 | -35.5779 | -36.4528
25.9553 | -35.5783 | -36.4586
25.9614 | -35.5787 | -36.4644
25.9675 | -35.5791 | -36.4701
25.9736 | -35.5795 | -36.4762
25.9797 | -35.5797 | -36.482
25.9858 | -35.5801 | -36.4879
25.9919 | -35.5805 | -36.4939
25.998 | -35.5808 | -36.4997
26.0042 | -35.5811 | -36.505
26.0103 | -35.5816 | -36.5107
26.0164 | -35.582 | -36.5164
26.0225 | -35.5824 | -36.5221
26.0286 | -35.5828 | -36.5278
26.0347 | -35.5831 | -36.5341
26.0408 | -35.5836 | -36.5396
26.0469 | -35.5839 | -36.5457
26.053 | -35.5842 | -36.5516
26.0591 | -35.5846 | -36.5573
26.0652 | -35.585 | -36.5627
26.0713 | -35.5854 | -36.5685
26.0774 | -35.5857 | -36.5741
26.0835 | -35.5861 | -36.5797
26.0896 | -35.5865 | -36.5856
26.0957 | -35.5869 | -36.5914
26.1018 | -35.5873 | -36.5969
26.1079 | -35.5877 | -36.6026
26.114 | -35.5879 | -36.6084
26.1201 | -35.5884 | -36.6141
26.1262 | -35.5888 | -36.6202
26.1323 | -35.5892 | -36.6259
26.1384 | -35.5896 | -36.6317
26.1445 | -35.5899 | -36.6374
26.1507 | -35.5903 | -36.6433
26.1568 | -35.5907 | -36.6487
26.1629 | -35.5911 | -36.6545
26.169 | -35.5914 | -36.6603
26.1751 | -35.5917 | -36.666
26.1812 | -35.5921 | -36.6717
26.1873 | -35.5925 | -36.6773
26.1934 | -35.5928 | -36.6829
26.1995 | -35.5932 | -36.6886
26.2056 | -35.5937 | -36.6943
26.2117 | -35.5941 | -36.6999
26.2178 | -35.5944 | -36.7057
26.2239 | -35.5947 | -36.7111
26.23 | -35.595 | -36.7169
26.2361 | -35.5954 | -36.7225
26.2422 | -35.5959 | -36.7285
26.2483 | -35.5963 | -36.7344
26.2544 | -35.5966 | -36.7398
26.2605 | -35.597 | -36.7455
26.2666 | -35.5973 | -36.7515
26.2727 | -35.5977 | -36.7574
26.2788 | -35.5981 | -36.7629
26.2849 | -35.5985 | -36.7686
26.2911 | -35.5989 | -36.7747
26.2972 | -35.5993 | -36.7805
26.3033 | -35.5997 | -36.786
26.3094 | -35.6 | -36.7918
26.3155 | -35.6003 | -36.7977
26.3216 | -35.6007 | -36.8037
26.3277 | -35.6011 | -36.8097
26.3338 | -35.6014 | -36.8152
26.3399 | -35.6018 | -36.821
26.346 | -35.6022 | -36.8265
26.3521 | -35.6025 | -36.8319
26.3582 | -35.6029 | -36.8376
26.3643 | -35.6032 | -36.8431
26.3704 | -35.6037 | -36.849
26.3765 | -35.604 | -36.8547
26.3826 | -35.6044 | -36.8604
26.3887 | -35.6048 | -36.8662
26.3948 | -35.6052 | -36.8719
26.4009 | -35.6055 | -36.8777
26.407 | -35.6059 | -36.8834
26.4131 | -35.6063 | -36.8892
26.4192 | -35.6066 | -36.8948
26.4253 | -35.6069 | -36.9007
26.4314 | -35.6073 | -36.9067
26.4376 | -35.6077 | -36.9122
26.4437 | -35.608 | -36.9178
26.4498 | -35.6084 | -36.9238
26.4559 | -35.6088 | -36.9293
26.462 | -35.6091 | -36.9347
26.4681 | -35.6095 | -36.9406
26.4742 | -35.6098 | -36.9466
26.4803 | -35.6101 | -36.9524
26.4864 | -35.6105 | -36.9581
26.4925 | -35.6109 | -36.9634
26.4986 | -35.6113 | -36.9694
26.5047 | -35.6116 | -36.9748
26.5108 | -35.612 | -36.9803
26.5169 | -35.6124 | -36.9857
26.523 | -35.6127 | -36.9914
26.5291 | -35.6131 | -36.9968
26.5352 | -35.6134 | -37.0026
26.5413 | -35.6137 | -37.0083
26.5474 | -35.6141 | -37.0141
26.5535 | -35.6145 | -37.0198
26.5596 | -35.6148 | -37.0256
26.5657 | -35.6152 | -37.0317
26.5718 | -35.6156 | -37.0373
26.578 | -35.616 | -37.0427
26.5841 | -35.6164 | -37.0479
26.5902 | -35.6167 | -37.0534
26.5963 | -35.6171 | -37.0589
26.6024 | -35.6175 | -37.0648
26.6085 | -35.6178 | -37.0706
26.6146 | -35.6181 | -37.0764
26.6207 | -35.6185 | -37.0818
26.6268 | -35.6189 | -37.0874
26.6329 | -35.6193 | -37.0929
26.639 | -35.6196 | -37.0988
26.6451 | -35.62 | -37.1045
26.6512 | -35.6204 | -37.11
26.6573 | -35.6208 | -37.1152
26.6634 | -35.6212 | -37.1208
26.6695 | -35.6216 | -37.1263
26.6756 | -35.622 | -37.1324
26.6817 | -35.6223 | -37.1381
26.6878 | -35.6227 | -37.1441
26.6939 | -35.6231 | -37.1499
26.7 | -35.6234 | -37.1554
26.7061 | -35.6239 | -37.1611
26.7122 | -35.6242 | -37.1668
26.7183 | -35.6246 | -37.1725
26.7245 | -35.625 | -37.1782
26.7306 | -35.6253 | -37.1841
26.7367 | -35.6257 | -37.1898
26.7428 | -35.626 | -37.1956
26.7489 | -35.6264 | -37.2013
26.755 | -35.6267 | -37.2071
26.7611 | -35.6271 | -37.2126
26.7672 | -35.6274 | -37.2183
26.7733 | -35.6278 | -37.2242
26.7794 | -35.6282 | -37.2299
26.7855 | -35.6285 | -37.2354
26.7916 | -35.6289 | -37.2409
26.7977 | -35.6292 | -37.2466
26.8038 | -35.6296 | -37.2515
26.8099 | -35.6299 | -37.2571
26.816 | -35.6303 | -37.2625
26.8221 | -35.6307 | -37.2686
26.8282 | -35.631 | -37.2744
26.8343 | -35.6314 | -37.2798
26.8404 | -35.6317 | -37.2852
26.8465 | -35.632 | -37.2905
26.8526 | -35.6324 | -37.2957
26.8587 | -35.6328 | -37.3012
26.8649 | -35.6332 | -37.3067
26.871 | -35.6336 | -37.3124
26.8771 | -35.634 | -37.3179
26.8832 | -35.6344 | -37.3237
26.8893 | -35.6347 | -37.3295
26.8954 | -35.635 | -37.3354
26.9015 | -35.6354 | -37.341
26.9076 | -35.6356 | -37.3465
26.9137 | -35.636 | -37.3518
26.9198 | -35.6364 | -37.3575
26.9259 | -35.6367 | -37.3631
26.932 | -35.6371 | -37.3688
26.9381 | -35.6375 | -37.3743
26.9442 | -35.6378 | -37.3802
26.9503 | -35.6381 | -37.3859
26.9564 | -35.6385 | -37.3915
26.9625 | -35.6389 | -37.397
26.9686 | -35.6392 | -37.4024
26.9747 | -35.6396 | -37.4081
26.9808 | -35.64 | -37.4136
26.9869 | -35.6403 | -37.4189
26.993 | -35.6407 | -37.4246
26.9991 | -35.6411 | -37.4302
27.0052 | -35.6414 | -37.4359
27.0114 | -35.6418 | -37.4414
27.0175 | -35.6421 | -37.4465
27.0236 | -35.6424 | -37.452
27.0297 | -35.6428 | -37.4579
27.0358 | -35.6431 | -37.4636
27.0419 | -35.6434 | -37.4691
27.048 | -35.6438 | -37.4746
27.0541 | -35.6442 | -37.4798
27.0602 | -35.6445 | -37.4854
27.0663 | -35.6448 | -37.4913
27.0724 | -35.6452 | -37.4971
27.0785 | -35.6455 | -37.5031
27.0846 | -35.6458 | -37.5086
27.0907 | -35.6461 | -37.5142
27.0968 | -35.6465 | -37.5199
27.1029 | -35.6469 | -37.5253
27.109 | -35.6472 | -37.5309
27.1151 | -35.6475 | -37.5364
27.1212 | -35.6478 | -37.542
27.1273 | -35.6482 | -37.5477
27.1334 | -35.6486 | -37.5534
27.1395 | -35.6489 | -37.5591
27.1456 | -35.6493 | -37.5648
27.1518 | -35.6496 | -37.5701
27.1579 | -35.6501 | -37.575
27.164 | -35.6504 | -37.5813
27.1701 | -35.6508 | -37.5864
27.1762 | -35.6511 | -37.5919
27.1823 | -35.6514 | -37.5973
27.1884 | -35.6518 | -37.6027
27.1945 | -35.6521 | -37.6086
27.2006 | -35.6525 | -37.6142
27.2067 | -35.6528 | -37.62
27.2128 | -35.6531 | -37.6255
27.2189 | -35.6534 | -37.6311
27.225 | -35.6538 | -37.6367
27.2311 | -35.6542 | -37.6418
27.2372 | -35.6545 | -37.6471
27.2433 | -35.6549 | -37.6527
27.2494 | -35.6552 | -37.6583
27.2555 | -35.6556 | -37.6641
27.2616 | -35.6559 | -37.6692
27.2677 | -35.6562 | -37.6746
27.2738 | -35.6565 | -37.6801
27.2799 | -35.6569 | -37.6856
27.286 | -35.6573 | -37.6908
27.2921 | -35.6577 | -37.6965
27.2983 | -35.658 | -37.7017
27.3044 | -35.6583 | -37.7072
27.3105 | -35.6587 | -37.7124
27.3166 | -35.6591 | -37.7178
27.3227 | -35.6594 | -37.7229
27.3288 | -35.6598 | -37.7287
27.3349 | -35.6601 | -37.7346
27.341 | -35.6604 | -37.7398
27.3471 | -35.6608 | -37.7455
27.3532 | -35.6611 | -37.7513
27.3593 | -35.6615 | -37.7568
27.3654 | -35.6619 | -37.7622
27.3715 | -35.6622 | -37.7673
27.3776 | -35.6625 | -37.7728
27.3837 | -35.6629 | -37.7784
27.3898 | -35.6632 | -37.7838
27.3959 | -35.6636 | -37.7892
27.402 | -35.664 | -37.7946
27.4081 | -35.6643 | -37.7998
27.4142 | -35.6647 | -37.8057
27.4203 | -35.665 | -37.8116
27.4264 | -35.6653 | -37.817
27.4325 | -35.6657 | -37.8224
27.4387 | -35.6659 | -37.8282
27.4448 | -35.6663 | -37.834
27.4509 | -35.6666 | -37.8395
27.457 | -35.667 | -37.8447
27.4631 | -35.6673 | -37.8507
27.4692 | -35.6676 | -37.8563
27.4753 | -35.668 | -37.8614
27.4814 | -35.6683 | -37.8667
27.4875 | -35.6687 | -37.8719
27.4936 | -35.6691 | -37.8772
27.4997 | -35.6693 | -37.8829
27.5058 | -35.6696 | -37.8878
27.5119 | -35.67 | -37.8932
27.518 | -35.6703 | -37.8985
27.5241 | -35.6707 | -37.9041
27.5302 | -35.671 | -37.9099
27.5363 | -35.6713 | -37.9151
27.5424 | -35.6717 | -37.9201
27.5485 | -35.6721 | -37.9253
27.5546 | -35.6725 | -37.9308
27.5607 | -35.6727 | -37.9362
27.5668 | -35.673 | -37.9417
27.5729 | -35.6733 | -37.9469
27.5791 | -35.6736 | -37.9521
27.5852 | -35.674 | -37.9576
27.5913 | -35.6744 | -37.9629
27.5974 | -35.6747 | -37.9685
27.6035 | -35.675 | -37.9738
27.6096 | -35.6754 | -37.9794
27.6157 | -35.6756 | -37.9845
27.6218 | -35.676 | -37.9901
27.6279 | -35.6764 | -37.9954
27.634 | -35.6767 | -38.0007
27.6401 | -35.6771 | -38.0062
27.6462 | -35.6774 | -38.0112
27.6523 | -35.6777 | -38.0161
27.6584 | -35.6781 | -38.0212
27.6645 | -35.6785 | -38.0265
27.6706 | -35.6789 | -38.0315
27.6767 | -35.6792 | -38.0367
27.6828 | -35.6796 | -38.042
27.6889 | -35.6799 | -38.0477
27.695 | -35.6802 | -38.0534
27.7011 | -35.6806 | -38.0589
27.7072 | -35.681 | -38.0643
27.7133 | -35.6813 | -38.0692
27.7194 | -35.6817 | -38.0749
27.7256 | -35.682 | -38.0803
27.7317 | -35.6823 | -38.0858
27.7378 | -35.6826 | -38.0918
27.7439 | -35.6829 | -38.0976
27.75 | -35.6832 | -38.1028
27.7561 | -35.6836 | -38.1088
27.7622 | -35.6839 | -38.114
27.7683 | -35.6843 | -38.1192
27.7744 | -35.6845 | -38.1241
27.7805 | -35.6848 | -38.1297
27.7866 | -35.6852 | -38.1355
27.7927 | -35.6855 | -38.1411
27.7988 | -35.6859 | -38.1461
27.8049 | -35.6862 | -38.1516
27.811 | -35.6866 | -38.1568
27.8171 | -35.6869 | -38.1619
27.8232 | -35.6872 | -38.167
27.8293 | -35.6875 | -38.172
27.8354 | -35.6879 | -38.178
27.8415 | -35.6882 | -38.1827
27.8476 | -35.6886 | -38.1878
27.8537 | -35.6889 | -38.193
27.8598 | -35.6891 | -38.1983
27.866 | -35.6894 | -38.2035
27.8721 | -35.6897 | -38.2091
27.8782 | -35.69 | -38.2143
27.8843 | -35.6903 | -38.2194
27.8904 | -35.6906 | -38.2254
27.8965 | -35.691 | -38.2305
27.9026 | -35.6914 | -38.2366
27.9087 | -35.6919 | -38.2423
27.9148 | -35.6922 | -38.2478
27.9209 | -35.6925 | -38.2531
27.927 | -35.6928 | -38.259
27.9331 | -35.6931 | -38.2642
27.9392 | -35.6934 | -38.2696
27.9453 | -35.6937 | -38.2746
27.9514 | -35.6941 | -38.2799
27.9575 | -35.6945 | -38.2857
27.9636 | -35.6948 | -38.2914
27.9697 | -35.6951 | -38.2972
27.9758 | -35.6953 | -38.3025
27.9819 | -35.6956 | -38.3078
27.988 | -35.696 | -38.3134
27.9941 | -35.6963 | -38.3184
28.0002 | -35.6965 | -38.3231
28.0063 | -35.6969 | -38.3286
28.0125 | -35.6971 | -38.3345
28.0186 | -35.6975 | -38.3401
28.0247 | -35.698 | -38.3453
28.0308 | -35.6982 | -38.3511
28.0369 | -35.6986 | -38.3559
28.043 | -35.6988 | -38.3614
28.0491 | -35.699 | -38.366
28.0552 | -35.6993 | -38.372
28.0613 | -35.6997 | -38.3767
28.0674 | -35.6998 | -38.3817
28.0735 | -35.7002 | -38.3862

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.4488 | -42.5101
21.4908 | -35.4994 | -35.2922
16.4948 | -34.8804 | -27.8308
12.6602 | -34.5225 | -21.506
9.7171 | -34.3118 | -16.5539
7.4582 | -34.1591 | -12.5608
5.7244 | -34.0537 | -9.3198
4.3936 | -33.9916 | -6.714
3.3722 | -33.9724 | -4.6577
2.5883 | -33.9664 | -3.101
1.9866 | -33.9451 | -1.8578
1.5248 | -33.9517 | -0.86081
1.1703 | -33.9586 | -0.20944
1 | -33.9692 | 0.11244
0.89822 | -33.9464 | 0.38227
0.68941 | -33.9365 | 0.96472
0.52919 | -33.9402 | 1.4107
0.40612 | -33.9413 | 2.017
0.31171 | -33.9277 | 2.4698
0.23928 | -33.9321 | 3.1885
0.18358 | -33.9286 | 3.6068
0.14091 | -33.9421 | 4.4842
0.10819 | -33.9644 | 5.6382
0.083017 | -33.9852 | 7.1094
0.063705 | -34.0359 | 8.9794
0.048923 | -34.1102 | 11.4664
0.037479 | -34.2072 | 14.6555
0.028801 | -34.3869 | 18.6867
0.022078 | -34.6939 | 24.0654
0.017023 | -35.1831 | 30.4911
0.013018 | -35.9755 | 38.1862
0.010014 | -37.0928 | 46.0999

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003658 | -18.1379
0.007316 | -34.7096
0.010974 | -37.6341
0.014633 | -39.3221
0.018291 | -40.7848
0.021949 | -43.4838
0.025607 | -45.9068
0.029265 | -46.66
0.032923 | -47.8901
0.036581 | -48.7368
0.04024 | -49.5441
0.043898 | -50.687
0.047556 | -51.4133
0.051214 | -52.4524
0.054872 | -53.8194
0.05853 | -54.3823
0.062188 | -55.9267
0.065847 | -56.5432
0.069505 | -55.3629
0.073163 | -55.5876
0.076821 | -57.1573
0.080479 | -58.8834
0.084137 | -59.5009
0.087796 | -58.648
0.091454 | -59.1955
0.095112 | -59.7677
0.09877 | -60.3077
0.10243 | -61.0691
0.10609 | -62.5368
0.10974 | -63.7747
0.1134 | -63.4043
0.11706 | -63.2934
0.12072 | -63.1861
0.12438 | -63.7082
0.12804 | -64.0238
0.13169 | -63.5328
0.13535 | -65.0157
0.13901 | -66.7458
0.14267 | -66.8652
0.14633 | -65.955
0.14998 | -67.4092
0.15364 | -66.7843
0.1573 | -66.3108
0.16096 | -67.0832
0.16462 | -67.8475
0.16828 | -68.6232
0.17193 | -68.106
0.17559 | -68.0236
0.17925 | -68.1151
0.18291 | -68.7232
0.18657 | -69.6337
0.19022 | -69.2931
0.19388 | -69.1858
0.19754 | -69.0644
0.2012 | -69.4719
0.20486 | -69.0596
0.20851 | -68.7541
0.21217 | -69.6181
0.21583 | -70.9099
0.21949 | -71.6921
0.22315 | -71.9036
0.22681 | -71.3096
0.23046 | -71.7659
0.23412 | -71.6
0.23778 | -72.381
0.24144 | -72.4976
0.2451 | -72.4435
0.24875 | -72.8856
0.25241 | -73.1023
0.25607 | -73.1913
0.25973 | -71.9694
0.26339 | -71.3331
0.26704 | -71.8349
0.2707 | -73.2254
0.27436 | -72.9893
0.27802 | -73.1019
0.28168 | -73.7413
0.28534 | -74.5962
0.28899 | -75.6685
0.29265 | -74.8979
0.29631 | -74.8053
0.29997 | -73.9423
0.30363 | -74.1916
0.30728 | -74.1242
0.31094 | -74.2463
0.3146 | -75.1771
0.31826 | -75.5823
0.32192 | -76.5727
0.32558 | -77.2777
0.32923 | -76.1245
0.33289 | -75.5511
0.33655 | -76.1421
0.34021 | -75.5018
0.34387 | -75.4248
0.34752 | -76.1416
0.35118 | -75.6753
0.35484 | -75.4417
0.3585 | -76.1805
0.36216 | -76.9187
0.36582 | -78.3608
0.36947 | -77.2938
0.37313 | -76.9025
0.37679 | -76.769
0.38045 | -77.5414
0.3841 | -78.028
0.38776 | -78.2692
0.39142 | -78.7448
0.39508 | -78.1613
0.39874 | -77.1682
0.4024 | -77.2508
0.40605 | -78.1208
0.40971 | -78.9371
0.41337 | -78.2964
0.41703 | -78.0728
0.42069 | -77.6857
0.42435 | -77.9868
0.428 | -78.4878
0.43166 | -77.748
0.43532 | -78.0541
0.43898 | -77.7759
0.44264 | -78.3381
0.44629 | -78.0915
0.44995 | -77.6144
0.45361 | -78.1502
0.45727 | -78.5184
0.46093 | -77.3731
0.46459 | -78.8093
0.46824 | -80.1268
0.4719 | -80.1445
0.47556 | -79.6771
0.47922 | -78.5789
0.48287 | -79.1299
0.48653 | -79.5297
0.49019 | -79.3777
0.49385 | -79.0614
0.49751 | -78.9403
0.50117 | -78.3148
0.50482 | -78.3697
0.50848 | -79.5061
0.51214 | -80.2655
0.5158 | -79.9423
0.51946 | -80.7378
0.52311 | -81.8079
0.52677 | -80.8769
0.53043 | -80.3616
0.53409 | -80.2023
0.53775 | -80.1091
0.54141 | -80.3294
0.54506 | -80.9291
0.54872 | -81.4403
0.55238 | -81.432
0.55604 | -81.1419
0.5597 | -81.16
0.56336 | -80.9962
0.56701 | -80.4833
0.57067 | -80.4692
0.57433 | -80.702
0.57799 | -81.0218
0.58164 | -81.7535
0.5853 | -82.297
0.58896 | -82.4178
0.59262 | -82.2179
0.59628 | -81.8779
0.59994 | -81.882
0.60359 | -82.4232
0.60725 | -82.794
0.61091 | -82.6099
0.61457 | -82.0934
0.61823 | -82.0779
0.62189 | -82.6442
0.62554 | -83.1052
0.6292 | -83.2102
0.63286 | -82.971
0.63652 | -82.4909
0.64018 | -82.4231
0.64383 | -82.5805
0.64749 | -82.4786
0.65115 | -82.2178
0.65481 | -82.0001
0.65847 | -81.9872
0.66212 | -82.2013
0.66578 | -82.7046
0.66944 | -83.4559
0.6731 | -83.8749
0.67676 | -83.8606
0.68041 | -83.7012
0.68407 | -83.5931
0.68773 | -83.507
0.69139 | -83.7846
0.69505 | -84.12
0.69871 | -83.9465
0.70236 | -83.5272
0.70602 | -83.7412
0.70968 | -84.2436
0.71334 | -84.219
0.717 | -84.2333
0.72066 | -84.3381
0.72431 | -83.955
0.72797 | -83.5447
0.73163 | -83.3874
0.73529 | -83.4528
0.73895 | -83.6657
0.7426 | -83.7959
0.74626 | -83.6417
0.74992 | -83.624
0.75358 | -83.7603
0.75724 | -83.3824
0.76089 | -83.2649
0.76455 | -83.998
0.76821 | -84.6723
0.77187 | -84.6896
0.77553 | -84.3355
0.77919 | -83.7521
0.78284 | -83.5397
0.7865 | -83.8518
0.79016 | -83.7996
0.79382 | -83.5375
0.79748 | -83.5598
0.80113 | -84.0182
0.80479 | -84.5945
0.80845 | -84.8412
0.81211 | -84.7368
0.81577 | -84.3866
0.81942 | -84.0949
0.82308 | -84.0934
0.82674 | -84.3728
0.8304 | -85.0878
0.83406 | -85.6748
0.83771 | -85.8925
0.84137 | -86.0515
0.84503 | -86.073
0.84869 | -85.8399
0.85235 | -85.1747
0.85601 | -85.0119
0.85966 | -85.4368
0.86332 | -85.4453
0.86698 | -85.2094
0.87064 | -85.5086
0.8743 | -86.2306
0.87796 | -86.7374
0.88161 | -86.4209
0.88527 | -85.5965
0.88893 | -85.1195
0.89259 | -85.5625
0.89625 | -85.9326
0.8999 | -85.6263
0.90356 | -85.1936
0.90722 | -85.0539
0.91088 | -84.6646
0.91454 | -84.4977
0.91819 | -85.0939
0.92185 | -85.4116
0.92551 | -85.1824
0.92917 | -84.9475
0.93283 | -85.1368
0.93649 | -85.5201
0.94014 | -85.9544
0.9438 | -86.2935
0.94746 | -86.2662
0.95112 | -86.2089
0.95478 | -86.5061
0.95843 | -86.5509
0.96209 | -86.1652
0.96575 | -86.1572
0.96941 | -86.4414
0.97307 | -86.2359
0.97672 | -85.5604
0.98038 | -85.4619
0.98404 | -86.1459
0.9877 | -87.0337
0.99136 | -87.4708
0.99502 | -87.1524
0.99867 | -86.4584
1.0023 | -86.0483
1.006 | -86.1042
1.0096 | -86.2852
1.0133 | -86.2344
1.017 | -86.1398
1.0206 | -86.2759
1.0243 | -86.7417
1.0279 | -86.8806
1.0316 | -86.5151
1.0353 | -86.6482
1.0389 | -87.1683
1.0426 | -87.6888
1.0462 | -87.6244
1.0499 | -86.8289
1.0535 | -86.6159
1.0572 | -86.6926
1.0609 | -86.9784
1.0645 | -87.1803
1.0682 | -87.2395
1.0718 | -87.319
1.0755 | -87.4133
1.0792 | -87.3731
1.0828 | -87.1991
1.0865 | -87.0477
1.0901 | -87.0947
1.0938 | -87.2172
1.0974 | -87.0036
1.1011 | -86.4524
1.1048 | -86.1817
1.1084 | -86.5937
1.1121 | -87.481
1.1157 | -88.2055
1.1194 | -88.3238
1.1231 | -87.9732
1.1267 | -87.5624
1.1304 | -87.369
1.134 | -87.4666
1.1377 | -87.729
1.1413 | -87.9947
1.145 | -88.2221
1.1487 | -88.3904
1.1523 | -88.3469
1.156 | -88.0053
1.1596 | -87.6129
1.1633 | -87.4395
1.1669 | -87.4967
1.1706 | -87.7249
1.1743 | -88.0392
1.1779 | -88.3753
1.1816 | -88.6848
1.1852 | -88.9057
1.1889 | -89.0799
1.1926 | -89.1583
1.1962 | -89.0313
1.1999 | -88.9162
1.2035 | -88.9904
1.2072 | -89.1931
1.2108 | -89.3596
1.2145 | -89.2541
1.2182 | -88.8337
1.2218 | -88.4002
1.2255 | -88.1706
1.2291 | -88.0694
1.2328 | -87.9007
1.2365 | -87.6991
1.2401 | -87.7275
1.2438 | -88.0893
1.2474 | -88.4842
1.2511 | -88.5106
1.2547 | -88.2502
1.2584 | -88.0545
1.2621 | -88.0446
1.2657 | -88.0941
1.2694 | -88.1557
1.273 | -88.3616
1.2767 | -88.7349
1.2804 | -89.1764
1.284 | -89.6284
1.2877 | -89.9946
1.2913 | -90.0566
1.295 | -89.7738
1.2986 | -89.4316
1.3023 | -89.1601
1.306 | -88.9207
1.3096 | -88.7868
1.3133 | -88.8254
1.3169 | -88.9487
1.3206 | -89.0834
1.3242 | -89.2541
1.3279 | -89.5019
1.3316 | -89.813
1.3352 | -90.1321
1.3389 | -90.2896
1.3425 | -90.2489
1.3462 | -90.1784
1.3499 | -90.1205
1.3535 | -90.0305
1.3572 | -89.9352
1.3608 | -89.7846
1.3645 | -89.6641
1.3681 | -89.6756
1.3718 | -89.6912
1.3755 | -89.6784
1.3791 | -89.7327
1.3828 | -89.7619
1.3864 | -89.6408
1.3901 | -89.3328
1.3938 | -88.9576
1.3974 | -88.7463
1.4011 | -88.8282
1.4047 | -89.1556
1.4084 | -89.5385
1.412 | -89.8559
1.4157 | -90.1801
1.4194 | -90.4407
1.423 | -90.292
1.4267 | -89.7712
1.4303 | -89.49
1.434 | -89.6447
1.4377 | -89.8752
1.4413 | -89.9702
1.445 | -90.0023
1.4486 | -89.9733
1.4523 | -89.8842
1.4559 | -89.8488
1.4596 | -89.9539
1.4633 | -90.0926
1.4669 | -90.1732
1.4706 | -90.3031
1.4742 | -90.5364
1.4779 | -90.7391
1.4815 | -90.8534
1.4852 | -90.9204
1.4889 | -90.8728
1.4925 | -90.6395
1.4962 | -90.3188
1.4998 | -90.1442
1.5035 | -90.1589
1.5072 | -90.3094
1.5108 | -90.686
1.5145 | -91.2542
1.5181 | -91.7077
1.5218 | -91.8291
1.5254 | -91.7113
1.5291 | -91.4223
1.5328 | -90.9856
1.5364 | -90.7032
1.5401 | -90.7308
1.5437 | -90.8449
1.5474 | -90.8324
1.5511 | -90.762
1.5547 | -90.7721
1.5584 | -90.8378
1.562 | -90.7371
1.5657 | -90.3187
1.5693 | -89.9219
1.573 | -89.9449
1.5767 | -90.3036
1.5803 | -90.5119
1.584 | -90.6764
1.5876 | -90.7946
1.5913 | -90.8811
1.595 | -90.9073
1.5986 | -90.8659
1.6023 | -90.8199
1.6059 | -90.8317
1.6096 | -90.9043
1.6132 | -90.9856
1.6169 | -91.0009
1.6206 | -90.9487
1.6242 | -90.9619
1.6279 | -91.1547
1.6315 | -91.5125
1.6352 | -91.9256
1.6388 | -92.2477
1.6425 | -92.3631
1.6462 | -92.2299
1.6498 | -91.9622
1.6535 | -91.6899
1.6571 | -91.4821
1.6608 | -91.3378
1.6645 | -91.2546
1.6681 | -91.2522
1.6718 | -91.2979
1.6754 | -91.3153
1.6791 | -91.3167
1.6827 | -91.3631
1.6864 | -91.4959
1.6901 | -91.6319
1.6937 | -91.6362
1.6974 | -91.4851
1.701 | -91.2802
1.7047 | -91.1342
1.7084 | -91.1188
1.712 | -91.2745
1.7157 | -91.5799
1.7193 | -91.9
1.723 | -92.106
1.7266 | -92.1367
1.7303 | -91.988
1.734 | -91.7138
1.7376 | -91.3985
1.7413 | -91.1316
1.7449 | -90.9955
1.7486 | -91.0443
1.7523 | -91.2588
1.7559 | -91.5691
1.7596 | -91.9146
1.7632 | -92.2559
1.7669 | -92.5464
1.7705 | -92.7561
1.7742 | -92.8728
1.7779 | -92.8866
1.7815 | -92.7357
1.7852 | -92.4046
1.7888 | -91.985
1.7925 | -91.6733
1.7961 | -91.6084
1.7998 | -91.8366
1.8035 | -92.2676
1.8071 | -92.6957
1.8108 | -92.8908
1.8144 | -92.8286
1.8181 | -92.6868
1.8218 | -92.6359
1.8254 | -92.7056
1.8291 | -92.8536
1.8327 | -92.9825
1.8364 | -93.0207
1.84 | -92.9713
1.8437 | -92.9519
1.8474 | -93.0387
1.851 | -93.1842
1.8547 | -93.2769
1.8583 | -93.2539
1.862 | -93.1056
1.8657 | -92.8331
1.8693 | -92.5275
1.873 | -92.3075
1.8766 | -92.2262
1.8803 | -92.2928
1.8839 | -92.4297
1.8876 | -92.5026
1.8913 | -92.4283
1.8949 | -92.2766
1.8986 | -92.1819
1.9022 | -92.1881
1.9059 | -92.2622
1.9096 | -92.3476
1.9132 | -92.4205
1.9169 | -92.5131
1.9205 | -92.6106
1.9242 | -92.6335
1.9278 | -92.4977
1.9315 | -92.2428
1.9352 | -92.0274
1.9388 | -92.001
1.9425 | -92.2036
1.9461 | -92.5651
1.9498 | -92.9612
1.9534 | -93.262
1.9571 | -93.3602
1.9608 | -93.2539
1.9644 | -93.0457
1.9681 | -92.8326
1.9717 | -92.6783
1.9754 | -92.5796
1.9791 | -92.5057
1.9827 | -92.4478
1.9864 | -92.482
1.99 | -92.66
1.9937 | -92.8603
1.9973 | -92.8749
2.001 | -92.6859
2.0047 | -92.489
2.0083 | -92.4842
2.012 | -92.7024
2.0156 | -93.0402
2.0193 | -93.3615
2.023 | -93.6112
2.0266 | -93.7683
2.0303 | -93.7972
2.0339 | -93.7503
2.0376 | -93.7172
2.0412 | -93.6873
2.0449 | -93.6132
2.0486 | -93.5363
2.0522 | -93.5291
2.0559 | -93.5535
2.0595 | -93.5336
2.0632 | -93.4069
2.0669 | -93.2375
2.0705 | -93.1383
2.0742 | -93.1567
2.0778 | -93.2365
2.0815 | -93.2816
2.0851 | -93.2442
2.0888 | -93.1621
2.0925 | -93.0795
2.0961 | -93.0701
2.0998 | -93.1708
2.1034 | -93.3738
2.1071 | -93.5922
2.1107 | -93.6574
2.1144 | -93.7043
2.1181 | -93.6864
2.1217 | -93.6208
2.1254 | -93.5095
2.129 | -93.3979
2.1327 | -93.3295
2.1364 | -93.3542
2.14 | -93.4935
2.1437 | -93.7171
2.1473 | -93.9838
2.151 | -94.2404
2.1546 | -94.439
2.1583 | -94.5397
2.162 | -94.5292
2.1656 | -94.4281
2.1693 | -94.2604
2.1729 | -94.0714
2.1766 | -93.8798
2.1803 | -93.69
2.1839 | -93.5026
2.1876 | -93.2965
2.1912 | -93.0895
2.1949 | -92.9261
2.1985 | -92.86
2.2022 | -92.9221
2.2059 | -93.101
2.2095 | -93.3416
2.2132 | -93.5571
2.2168 | -93.6839
2.2205 | -93.7248
2.2242 | -93.7177
2.2278 | -93.7127
2.2315 | -93.7349
2.2351 | -93.7886
2.2388 | -93.856
2.2424 | -93.9161
2.2461 | -93.9191
2.2498 | -93.8638
2.2534 | -93.7784
2.2571 | -93.6992
2.2607 | -93.6401
2.2644 | -93.6037
2.2681 | -93.5978
2.2717 | -93.6468
2.2754 | -93.7792
2.279 | -93.9744
2.2827 | -94.1598
2.2863 | -94.2999
2.29 | -94.3805
2.2937 | -94.4523
2.2973 | -94.5469
2.301 | -94.6586
2.3046 | -94.7635
2.3083 | -94.8066
2.3119 | -94.7757
2.3156 | -94.6831
2.3193 | -94.5858
2.3229 | -94.5403
2.3266 | -94.5792
2.3302 | -94.6904
2.3339 | -94.7843
2.3376 | -94.7756
2.3412 | -94.6515
2.3449 | -94.4735
2.3485 | -94.303
2.3522 | -94.1843
2.3558 | -94.1431
2.3595 | -94.1742
2.3632 | -94.2587
2.3668 | -94.359
2.3705 | -94.4693
2.3741 | -94.5919
2.3778 | -94.6915
2.3815 | -94.716
2.3851 | -94.6545
2.3888 | -94.5543
2.3924 | -94.4621
2.3961 | -94.4193
2.3997 | -94.4435
2.4034 | -94.5001
2.4071 | -94.5516
2.4107 | -94.5707
2.4144 | -94.5697
2.418 | -94.576
2.4217 | -94.6046
2.4254 | -94.6392
2.429 | -94.6627
2.4327 | -94.649
2.4363 | -94.6101
2.44 | -94.543
2.4436 | -94.4496
2.4473 | -94.3304
2.451 | -94.2166
2.4546 | -94.1621
2.4583 | -94.1933
2.4619 | -94.2944
2.4656 | -94.4387
2.4692 | -94.5786
2.4729 | -94.6565
2.4766 | -94.6481
2.4802 | -94.6024
2.4839 | -94.581
2.4875 | -94.5915
2.4912 | -94.5929
2.4949 | -94.5723
2.4985 | -94.57
2.5022 | -94.6168
2.5058 | -94.7051
2.5095 | -94.7972
2.5131 | -94.8372
2.5168 | -94.8136
2.5205 | -94.7812
2.5241 | -94.7729
2.5278 | -94.7749
2.5314 | -94.7259
2.5351 | -94.5834
2.5388 | -94.3612
2.5424 | -94.1427
2.5461 | -94.0094
2.5497 | -93.9979
2.5534 | -94.0996
2.557 | -94.261
2.5607 | -94.4323
2.5644 | -94.6108
2.568 | -94.7581
2.5717 | -94.8344
2.5753 | -94.8483
2.579 | -94.8575
2.5827 | -94.887
2.5863 | -94.9186
2.59 | -94.9354
2.5936 | -94.9402
2.5973 | -94.9608
2.6009 | -95.0247
2.6046 | -95.1181
2.6083 | -95.2338
2.6119 | -95.3626
2.6156 | -95.4662
2.6192 | -95.518
2.6229 | -95.5159
2.6265 | -95.4705
2.6302 | -95.4201
2.6339 | -95.3789
2.6375 | -95.3795
2.6412 | -95.3698
2.6448 | -95.3532
2.6485 | -95.3229
2.6522 | -95.2624
2.6558 | -95.1647
2.6595 | -95.0685
2.6631 | -94.9975
2.6668 | -94.9514
2.6704 | -94.9391
2.6741 | -94.9489
2.6778 | -94.9817
2.6814 | -95.0216
2.6851 | -95.0573
2.6887 | -95.1084
2.6924 | -95.1804
2.6961 | -95.2748
2.6997 | -95.3854
2.7034 | -95.4921
2.707 | -95.572
2.7107 | -95.6143
2.7143 | -95.6297
2.718 | -95.6256
2.7217 | -95.6288
2.7253 | -95.673
2.729 | -95.7534
2.7326 | -95.8496
2.7363 | -95.91
2.74 | -95.9005
2.7436 | -95.8263
2.7473 | -95.7108
2.7509 | -95.6029
2.7546 | -95.5354
2.7582 | -95.5411
2.7619 | -95.6021
2.7656 | -95.6887
2.7692 | -95.7757
2.7729 | -95.8422
2.7765 | -95.8965
2.7802 | -95.9259
2.7838 | -95.9455
2.7875 | -95.9496
2.7912 | -95.9154
2.7948 | -95.8523
2.7985 | -95.7593
2.8021 | -95.6622
2.8058 | -95.581
2.8095 | -95.5485
2.8131 | -95.5659
2.8168 | -95.628
2.8204 | -95.7139
2.8241 | -95.7858
2.8277 | -95.8141
2.8314 | -95.7788
2.8351 | -95.7092
2.8387 | -95.6521
2.8424 | -95.6492
2.846 | -95.7327
2.8497 | -95.8859
2.8534 | -96.068
2.857 | -96.226
2.8607 | -96.313
2.8643 | -96.3031
2.868 | -96.2014
2.8716 | -96.0536
2.8753 | -95.8983
2.879 | -95.7692
2.8826 | -95.7053
2.8863 | -95.7278
2.8899 | -95.8212
2.8936 | -95.9341
2.8973 | -95.9958
2.9009 | -95.96
2.9046 | -95.8303
2.9082 | -95.6593
2.9119 | -95.4998
2.9155 | -95.396
2.9192 | -95.361
2.9229 | -95.4161
2.9265 | -95.5345
2.9302 | -95.6723
2.9338 | -95.7868
2.9375 | -95.8686
2.9411 | -95.9399
2.9448 | -96.0184
2.9485 | -96.11
2.9521 | -96.212
2.9558 | -96.3001
2.9594 | -96.3622
2.9631 | -96.4064
2.9668 | -96.4194
2.9704 | -96.388
2.9741 | -96.3334
2.9777 | -96.2565
2.9814 | -96.1565
2.985 | -96.0347
2.9887 | -95.8981
2.9924 | -95.7664
2.996 | -95.6612
2.9997 | -95.6074
3.0033 | -95.6171
3.007 | -95.674
3.0107 | -95.7448
3.0143 | -95.8087
3.018 | -95.8535
3.0216 | -95.878
3.0253 | -95.9113
3.0289 | -95.9851
3.0326 | -96.1166
3.0363 | -96.3042
3.0399 | -96.5298
3.0436 | -96.754
3.0472 | -96.9287
3.0509 | -97.0271
3.0546 | -97.0333
3.0582 | -96.9539
3.0619 | -96.8207
3.0655 | -96.6666
3.0692 | -96.5335
3.0728 | -96.4603
3.0765 | -96.4536
3.0802 | -96.5019
3.0838 | -96.5759
3.0875 | -96.6409
3.0911 | -96.6772
3.0948 | -96.6935
3.0985 | -96.7002
3.1021 | -96.7143
3.1058 | -96.7457
3.1094 | -96.7867
3.1131 | -96.8133
3.1167 | -96.8149
3.1204 | -96.7733
3.1241 | -96.7006
3.1277 | -96.5964
3.1314 | -96.4759
3.135 | -96.3553
3.1387 | -96.2578
3.1423 | -96.1975
3.146 | -96.1708
3.1497 | -96.1771
3.1533 | -96.2056
3.157 | -96.2585
3.1606 | -96.3221
3.1643 | -96.3358
3.168 | -96.3557
3.1716 | -96.3587
3.1753 | -96.3546
3.1789 | -96.3583
3.1826 | -96.3841
3.1862 | -96.4354
3.1899 | -96.5028
3.1936 | -96.563
3.1972 | -96.6074
3.2009 | -96.6207
3.2045 | -96.6084
3.2082 | -96.5707
3.2119 | -96.523
3.2155 | -96.4674
3.2192 | -96.4125
3.2228 | -96.3767
3.2265 | -96.3823
3.2301 | -96.4324
3.2338 | -96.5262
3.2375 | -96.6488
3.2411 | -96.7826
3.2448 | -96.9067
3.2484 | -97.0247
3.2521 | -97.1255
3.2557 | -97.1982
3.2594 | -97.242
3.2631 | -97.2495
3.2667 | -97.2214
3.2704 | -97.1684
3.274 | -97.0941
3.2777 | -96.996
3.2814 | -96.895
3.285 | -96.8159
3.2887 | -96.7656
3.2923 | -96.7491
3.296 | -96.7659
3.2996 | -96.8002
3.3033 | -96.8325
3.307 | -96.8693
3.3106 | -96.9138
3.3143 | -96.9719
3.3179 | -97.0441
3.3216 | -97.1165
3.3253 | -97.1714
3.3289 | -97.2111
3.3326 | -97.2306
3.3362 | -97.2051
3.3399 | -97.1382
3.3435 | -97.05
3.3472 | -96.9594
3.3509 | -96.8923
3.3545 | -96.8711
3.3582 | -96.9045
3.3618 | -96.9845
3.3655 | -97.084
3.3692 | -97.1916
3.3728 | -97.3029
3.3765 | -97.4083
3.3801 | -97.5069
3.3838 | -97.5831
3.3874 | -97.6351
3.3911 | -97.6488
3.3948 | -97.6299
3.3984 | -97.5942
3.4021 | -97.5411
3.4057 | -97.4812
3.4094 | -97.4176
3.4131 | -97.344
3.4167 | -97.2746
3.4204 | -97.2253
3.424 | -97.1874
3.4277 | -97.156
3.4313 | -97.1354
3.435 | -97.1245
3.4387 | -97.1241
3.4423 | -97.1386
3.446 | -97.17
3.4496 | -97.1951
3.4533 | -97.2031
3.4569 | -97.1889
3.4606 | -97.1522
3.4643 | -97.0982
3.4679 | -97.0435
3.4716 | -96.9921
3.4752 | -96.9621
3.4789 | -96.9632
3.4826 | -96.999
3.4862 | -97.0634
3.4899 | -97.1535
3.4935 | -97.2497
3.4972 | -97.3395
3.5008 | -97.4097
3.5045 | -97.4605
3.5082 | -97.5017
3.5118 | -97.5566
3.5155 | -97.6015
3.5191 | -97.6472
3.5228 | -97.6789
3.5265 | -97.6952
3.5301 | -97.688
3.5338 | -97.6599
3.5374 | -97.63
3.5411 | -97.6044
3.5447 | -97.5723
3.5484 | -97.5377
3.5521 | -97.4891
3.5557 | -97.4196
3.5594 | -97.3414
3.563 | -97.2566
3.5667 | -97.17
3.5704 | -97.1016
3.574 | -97.0445
3.5777 | -97.0151
3.5813 | -97.0174
3.585 | -97.0697
3.5886 | -97.17
3.5923 | -97.3098
3.596 | -97.4673
3.5996 | -97.6399
3.6033 | -97.8038
3.6069 | -97.9473
3.6106 | -98.059
3.6142 | -98.1102
3.6179 | -98.1032
3.6216 | -98.0415
3.6252 | -97.9442
3.6289 | -97.8501
3.6325 | -97.7823
3.6362 | -97.7594
3.6399 | -97.7676
3.6435 | -97.8108
3.6472 | -97.8555
3.6508 | -97.8853
3.6545 | -97.8839
3.6581 | -97.8659
3.6618 | -97.8401
3.6655 | -97.8305
3.6691 | -97.8494
3.6728 | -97.8896
3.6764 | -97.9388
3.6801 | -97.9806
3.6838 | -97.998
3.6874 | -97.9905
3.6911 | -97.9251
3.6947 | -97.8845
3.6984 | -97.8485
3.702 | -97.8162
3.7057 | -97.8062
3.7094 | -97.8247
3.713 | -97.8672
3.7167 | -97.9191
3.7203 | -97.9722
3.724 | -98.0173
3.7277 | -98.0435
3.7313 | -98.0529
3.735 | -98.0447
3.7386 | -98.0192
3.7423 | -97.982
3.7459 | -97.9311
3.7496 | -97.8885
3.7533 | -97.8489
3.7569 | -97.8188
3.7606 | -97.8027
3.7642 | -97.8022
3.7679 | -97.815
3.7715 | -97.8388
3.7752 | -97.8608
3.7789 | -97.8831
3.7825 | -97.8988
3.7862 | -97.9064
3.7898 | -97.9205
3.7935 | -97.9376
3.7972 | -97.9636
3.8008 | -97.9971
3.8045 | -98.0311
3.8081 | -98.0784
3.8118 | -98.1363
3.8154 | -98.211
3.8191 | -98.2974
3.8228 | -98.3819
3.8264 | -98.4603
3.8301 | -98.5007
3.8337 | -98.4884
3.8374 | -98.4322
3.8411 | -98.3488
3.8447 | -98.2574
3.8484 | -98.1827
3.852 | -98.1271
3.8557 | -98.1015
3.8593 | -98.1035
3.863 | -98.1188
3.8667 | -98.1298
3.8703 | -98.1215
3.874 | -98.0923
3.8776 | -98.0534
3.8813 | -98.0107
3.885 | -97.9765
3.8886 | -97.9567
3.8923 | -97.9624
3.8959 | -97.9951
3.8996 | -98.0432
3.9032 | -98.0974
3.9069 | -98.1516
3.9106 | -98.1958
3.9142 | -98.2412
3.9179 | -98.287
3.9215 | -98.3431
3.9252 | -98.4176
3.9288 | -98.515
3.9325 | -98.6331
3.9362 | -98.7537
3.9398 | -98.8577
3.9435 | -98.9293
3.9471 | -98.9537
3.9508 | -98.9296
3.9545 | -98.8602
3.9581 | -98.7549
3.9618 | -98.6464
3.9654 | -98.5414
3.9691 | -98.4475
3.9727 | -98.3788
3.9764 | -98.3269
3.9801 | -98.2973
3.9837 | -98.292
3.9874 | -98.3161
3.991 | -98.3632
3.9947 | -98.4214
3.9984 | -98.4837
4.002 | -98.5484
4.0057 | -98.6066
4.0093 | -98.6563
4.013 | -98.701
4.0166 | -98.754
4.0203 | -98.8006
4.024 | -98.8269
4.0276 | -98.8361
4.0313 | -98.8224
4.0349 | -98.7741
4.0386 | -98.7166
4.0423 | -98.6522
4.0459 | -98.5957
4.0496 | -98.5536
4.0532 | -98.5364
4.0569 | -98.547
4.0605 | -98.5824
4.0642 | -98.6375
4.0679 | -98.7068
4.0715 | -98.7774
4.0752 | -98.8255
4.0788 | -98.839
4.0825 | -98.8257
4.0861 | -98.7929
4.0898 | -98.7465
4.0935 | -98.6995
4.0971 | -98.6574
4.1008 | -98.6232
4.1044 | -98.6112
4.1081 | -98.6173
4.1118 | -98.6471
4.1154 | -98.6928
4.1191 | -98.7592
4.1227 | -98.8373
4.1264 | -98.9106
4.13 | -98.974
4.1337 | -99.0132
4.1374 | -99.0206
4.141 | -98.9912
4.1447 | -98.9393
4.1483 | -98.8797
4.152 | -98.8231
4.1557 | -98.7829
4.1593 | -98.7743
4.163 | -98.7841
4.1666 | -98.8071
4.1703 | -98.8336
4.1739 | -98.8556
4.1776 | -98.8738
4.1813 | -98.8866
4.1849 | -98.9086
4.1886 | -98.9425
4.1922 | -98.9863
4.1959 | -99.046
4.1996 | -99.1154
4.2032 | -99.1764
4.2069 | -99.2115
4.2105 | -99.2272
4.2142 | -99.2248
4.2178 | -99.2141
4.2215 | -99.1988
4.2252 | -99.1862
4.2288 | -99.1712
4.2325 | -99.1541
4.2361 | -99.1429
4.2398 | -99.1454
4.2435 | -99.1567
4.2471 | -99.1792
4.2508 | -99.2169
4.2544 | -99.2671
4.2581 | -99.3169
4.2617 | -99.3559
4.2654 | -99.3823
4.2691 | -99.3998
4.2727 | -99.4069
4.2764 | -99.3928
4.28 | -99.3618
4.2837 | -99.3097
4.2873 | -99.2482
4.291 | -99.181
4.2947 | -99.119
4.2983 | -99.0632
4.302 | -99.009
4.3056 | -98.9531
4.3093 | -98.9064
4.313 | -98.874
4.3166 | -98.8543
4.3203 | -98.8447
4.3239 | -98.8486
4.3276 | -98.8784
4.3312 | -98.9265
4.3349 | -98.9864
4.3386 | -99.0505
4.3422 | -99.1191
4.3459 | -99.1893
4.3495 | -99.2546
4.3532 | -99.314
4.3569 | -99.3652
4.3605 | -99.4087
4.3642 | -99.4328
4.3678 | -99.4363
4.3715 | -99.4258
4.3751 | -99.4052
4.3788 | -99.3736
4.3825 | -99.3545
4.3861 | -99.3527
4.3898 | -99.3626
4.3934 | -99.3896
4.3971 | -99.4188
4.4008 | -99.4462
4.4044 | -99.4707
4.4081 | -99.4828
4.4117 | -99.4834
4.4154 | -99.4674
4.419 | -99.4372
4.4227 | -99.3942
4.4264 | -99.3508
4.43 | -99.3235
4.4337 | -99.3194
4.4373 | -99.3383
4.441 | -99.3726
4.4446 | -99.4216
4.4483 | -99.4767
4.452 | -99.5263
4.4556 | -99.5728
4.4593 | -99.6144
4.4629 | -99.6373
4.4666 | -99.6449
4.4703 | -99.6346
4.4739 | -99.6088
4.4776 | -99.5778
4.4812 | -99.5413
4.4849 | -99.5066
4.4885 | -99.487
4.4922 | -99.4899
4.4959 | -99.5052
4.4995 | -99.5262
4.5032 | -99.5427
4.5068 | -99.5563
4.5105 | -99.5582
4.5142 | -99.5505
4.5178 | -99.5421
4.5215 | -99.5341
4.5251 | -99.5256
4.5288 | -99.5216
4.5324 | -99.5191
4.5361 | -99.5159
4.5398 | -99.5182
4.5434 | -99.5254
4.5471 | -99.5466
4.5507 | -99.5787
4.5544 | -99.6224
4.558 | -99.6564
4.5617 | -99.676
4.5654 | -99.6771
4.569 | -99.6545
4.5727 | -99.6154
4.5763 | -99.5729
4.58 | -99.5315
4.5837 | -99.4969
4.5873 | -99.4736
4.591 | -99.4742
4.5946 | -99.4919
4.5983 | -99.5267
4.6019 | -99.5765
4.6056 | -99.6357
4.6093 | -99.7014
4.6129 | -99.7733
4.6166 | -99.8392
4.6202 | -99.8945
4.6239 | -99.9418
4.6276 | -99.9782
4.6312 | -99.994
4.6349 | -99.9934
4.6385 | -99.978
4.6422 | -99.9457
4.6458 | -99.9006
4.6495 | -99.8399
4.6532 | -99.7633
4.6568 | -99.6767
4.6605 | -99.5754
4.6641 | -99.4811
4.6678 | -99.4009
4.6715 | -99.3464
4.6751 | -99.3219
4.6788 | -99.3278
4.6824 | -99.3609
4.6861 | -99.4246
4.6897 | -99.5011
4.6934 | -99.5688
4.6971 | -99.6306
4.7007 | -99.6778
4.7044 | -99.7126
4.708 | -99.7379
4.7117 | -99.7565
4.7153 | -99.7735
4.719 | -99.8072
4.7227 | -99.8404
4.7263 | -99.8763
4.73 | -99.9207
4.7336 | -99.9593
4.7373 | -99.9893
4.741 | -100.009
4.7446 | -100.0216
4.7483 | -100.0227
4.7519 | -100.0264
4.7556 | -100.0327
4.7592 | -100.0329
4.7629 | -100.0345
4.7666 | -100.0337
4.7702 | -100.0269
4.7739 | -100.0218
4.7775 | -100.0115
4.7812 | -99.9982
4.7849 | -99.9802
4.7885 | -99.9638
4.7922 | -99.9511
4.7958 | -99.9411
4.7995 | -99.9215
4.8031 | -99.8973
4.8068 | -99.8818
4.8105 | -99.8724
4.8141 | -99.8591
4.8178 | -99.8513
4.8214 | -99.849
4.8251 | -99.8515
4.8288 | -99.8654
4.8324 | -99.8686
4.8361 | -99.876
4.8397 | -99.8849
4.8434 | -99.8928
4.847 | -99.8944
4.8507 | -99.8902
4.8544 | -99.8938
4.858 | -99.9058
4.8617 | -99.9261
4.8653 | -99.9498
4.869 | -99.9787
4.8727 | -100.0123
4.8763 | -100.0538
4.88 | -100.087
4.8836 | -100.1008
4.8873 | -100.1082
4.8909 | -100.1149
4.8946 | -100.1129
4.8983 | -100.1117
4.9019 | -100.1125
4.9056 | -100.1083
4.9092 | -100.1162
4.9129 | -100.1303
4.9165 | -100.1474
4.9202 | -100.1597
4.9239 | -100.1643
4.9275 | -100.1642
4.9312 | -100.1445
4.9348 | -100.1181
4.9385 | -100.0934
4.9422 | -100.0683
4.9458 | -100.0446
4.9495 | -100.0284
4.9531 | -100.0094
4.9568 | -99.9963
4.9604 | -99.9972
4.9641 | -100.0046
4.9678 | -100.0153
4.9714 | -100.0342
4.9751 | -100.0681
4.9787 | -100.1258
4.9824 | -100.1974
4.9861 | -100.2879
4.9897 | -100.3878
4.9934 | -100.5014
4.997 | -100.6268
5.0007 | -100.753
5.0043 | -100.8628
5.008 | -100.9581
5.0117 | -101.0202
5.0153 | -101.03
5.019 | -100.9993
5.0226 | -100.9291
5.0263 | -100.8215
5.03 | -100.6925
5.0336 | -100.5583
5.0373 | -100.4209
5.0409 | -100.3009
5.0446 | -100.2129
5.0482 | -100.1519
5.0519 | -100.1215
5.0556 | -100.127
5.0592 | -100.1629
5.0629 | -100.213
5.0665 | -100.2807
5.0702 | -100.3531
5.0738 | -100.4223
5.0775 | -100.483
5.0812 | -100.5337
5.0848 | -100.563
5.0885 | -100.5768
5.0921 | -100.5911
5.0958 | -100.5862
5.0995 | -100.5574
5.1031 | -100.5255
5.1068 | -100.4916
5.1104 | -100.4426
5.1141 | -100.3947
5.1177 | -100.3453
5.1214 | -100.2985
5.1251 | -100.2593
5.1287 | -100.2205
5.1324 | -100.1871
5.136 | -100.1525
5.1397 | -100.1252
5.1434 | -100.1049
5.147 | -100.0827
5.1507 | -100.0593
5.1543 | -100.0291
5.158 | -99.9943
5.1616 | -99.9623
5.1653 | -99.9304
5.169 | -99.8992
5.1726 | -99.8873
5.1763 | -99.9017
5.1799 | -99.9345
5.1836 | -99.9939
5.1873 | -100.0826
5.1909 | -100.183
5.1946 | -100.2838
5.1982 | -100.3745
5.2019 | -100.4539
5.2055 | -100.5126
5.2092 | -100.5516
5.2129 | -100.5898
5.2165 | -100.6112
5.2202 | -100.6273
5.2238 | -100.6527
5.2275 | -100.6764
5.2311 | -100.7077
5.2348 | -100.7469
5.2385 | -100.7703
5.2421 | -100.7774
5.2458 | -100.769
5.2494 | -100.7382
5.2531 | -100.6912
5.2568 | -100.6372
5.2604 | -100.5852
5.2641 | -100.5409
5.2677 | -100.5076
5.2714 | -100.4945
5.275 | -100.5239
5.2787 | -100.5389
5.2824 | -100.5622
5.286 | -100.607
5.2897 | -100.6661
5.2933 | -100.7326
5.297 | -100.8061
5.3007 | -100.8783
5.3043 | -100.9485
5.308 | -101.0181
5.3116 | -101.0728
5.3153 | -101.1052
5.3189 | -101.1405
5.3226 | -101.1726
5.3263 | -101.1946
5.3299 | -101.2053
5.3336 | -101.2
5.3372 | -101.1727
5.3409 | -101.1345
5.3446 | -101.0958
5.3482 | -101.0619
5.3519 | -101.0251
5.3555 | -101.0011
5.3592 | -100.9898
5.3628 | -100.9772
5.3665 | -100.977
5.3702 | -100.9791
5.3738 | -100.9857
5.3775 | -100.9962
5.3811 | -101.0163
5.3848 | -101.0413
5.3884 | -101.0612
5.3921 | -101.0794
5.3958 | -101.1037
5.3994 | -101.1354
5.4031 | -101.1682
5.4067 | -101.2013
5.4104 | -101.2314
5.4141 | -101.2671
5.4177 | -101.3097
5.4214 | -101.3483
5.425 | -101.383
5.4287 | -101.4088
5.4323 | -101.4238
5.436 | -101.4417
5.4397 | -101.4554
5.4433 | -101.4688
5.447 | -101.483
5.4506 | -101.4897
5.4543 | -101.4962
5.458 | -101.5004
5.4616 | -101.4977
5.4653 | -101.4964
5.4689 | -101.4902
5.4726 | -101.4809
5.4762 | -101.4683
5.4799 | -101.447
5.4836 | -101.4161
5.4872 | -101.3775
5.4909 | -101.3357
5.4945 | -101.3029
5.4982 | -101.2742
5.5019 | -101.2603
5.5055 | -101.2704
5.5092 | -101.3003
5.5128 | -101.3522
5.5165 | -101.4166
5.5201 | -101.4846
5.5238 | -101.5573
5.5275 | -101.6176
5.5311 | -101.6644
5.5348 | -101.6931
5.5384 | -101.7034
5.5421 | -101.6894
5.5457 | -101.6613
5.5494 | -101.6216
5.5531 | -101.5754
5.5567 | -101.5225
5.5604 | -101.4612
5.564 | -101.3985
5.5677 | -101.3389
5.5714 | -101.2716
5.575 | -101.2024
5.5787 | -101.1296
5.5823 | -101.0582
5.586 | -100.9931
5.5896 | -100.9332
5.5933 | -100.8873
5.597 | -100.8594
5.6006 | -100.8507
5.6043 | -100.8494
5.6079 | -100.8562
5.6116 | -100.8863
5.6153 | -100.9307
5.6189 | -100.9837
5.6226 | -101.0474
5.6262 | -101.1194
5.6299 | -101.1885
5.6335 | -101.2483
5.6372 | -101.3108
5.6409 | -101.3675
5.6445 | -101.4153
5.6482 | -101.4599
5.6518 | -101.4913
5.6555 | -101.5229
5.6592 | -101.5463
5.6628 | -101.5537
5.6665 | -101.5458
5.6701 | -101.5216
5.6738 | -101.4929
5.6774 | -101.4543
5.6811 | -101.4091
5.6848 | -101.3594
5.6884 | -101.3214
5.6921 | -101.2838
5.6957 | -101.246
5.6994 | -101.2145
5.7031 | -101.1886
5.7067 | -101.1633
5.7104 | -101.1467
5.714 | -101.1411
5.7177 | -101.135
5.7213 | -101.1267
5.725 | -101.1255
5.7287 | -101.1278
5.7323 | -101.1262
5.736 | -101.1391
5.7396 | -101.1614
5.7433 | -101.1931
5.7469 | -101.2404
5.7506 | -101.2985
5.7543 | -101.3677
5.7579 | -101.4309
5.7616 | -101.4918
5.7652 | -101.5489
5.7689 | -101.5958
5.7726 | -101.6221
5.7762 | -101.6326
5.7799 | -101.6214
5.7835 | -101.5969
5.7872 | -101.5629
5.7908 | -101.5243
5.7945 | -101.4849
5.7982 | -101.4466
5.8018 | -101.4359
5.8055 | -101.4148
5.8091 | -101.3923
5.8128 | -101.3714
5.8165 | -101.3524
5.8201 | -101.3338
5.8238 | -101.308
5.8274 | -101.2783
5.8311 | -101.2471
5.8347 | -101.2199
5.8384 | -101.1923
5.8421 | -101.1628
5.8457 | -101.1491
5.8494 | -101.1376
5.853 | -101.1257
5.8567 | -101.1212
5.8604 | -101.1226
5.864 | -101.1307
5.8677 | -101.155
5.8713 | -101.1875
5.875 | -101.2278
5.8786 | -101.2762
5.8823 | -101.3248
5.886 | -101.3789
5.8896 | -101.4371
5.8933 | -101.5
5.8969 | -101.569
5.9006 | -101.6374
5.9042 | -101.6937
5.9079 | -101.7426
5.9116 | -101.7891
5.9152 | -101.8285
5.9189 | -101.857
5.9225 | -101.8796
5.9262 | -101.8873
5.9299 | -101.8812
5.9335 | -101.862
5.9372 | -101.8308
5.9408 | -101.7962
5.9445 | -101.7597
5.9481 | -101.7206
5.9518 | -101.683
5.9555 | -101.652
5.9591 | -101.618
5.9628 | -101.5808
5.9664 | -101.5475
5.9701 | -101.5195
5.9738 | -101.4955
5.9774 | -101.4652
5.9811 | -101.4289
5.9847 | -101.3937
5.9884 | -101.3639
5.992 | -101.341
5.9957 | -101.3301
5.9994 | -101.3205
6.003 | -101.3135
6.0067 | -101.3137
6.0103 | -101.3224
6.014 | -101.3349
6.0177 | -101.3509
6.0213 | -101.3729
6.025 | -101.3935
6.0286 | -101.4151
6.0323 | -101.4386
6.0359 | -101.4664
6.0396 | -101.4954
6.0433 | -101.5192
6.0469 | -101.5368
6.0506 | -101.5485
6.0542 | -101.5587
6.0579 | -101.5609
6.0615 | -101.5532
6.0652 | -101.5432
6.0689 | -101.54
6.0725 | -101.547
6.0762 | -101.558
6.0798 | -101.5765
6.0835 | -101.6086
6.0872 | -101.6452
6.0908 | -101.6887
6.0945 | -101.7339
6.0981 | -101.779
6.1018 | -101.8275
6.1054 | -101.8717
6.1091 | -101.9238
6.1128 | -101.9745
6.1164 | -102.0218
6.1201 | -102.0589
6.1237 | -102.0913
6.1274 | -102.1193
6.1311 | -102.1353
6.1347 | -102.1403
6.1384 | -102.1408
6.142 | -102.1367
6.1457 | -102.1249
6.1493 | -102.1094
6.153 | -102.0904
6.1567 | -102.0699
6.1603 | -102.0481
6.164 | -102.0246
6.1676 | -102.0059
6.1713 | -101.9853
6.175 | -101.9702
6.1786 | -101.9582
6.1823 | -101.9473
6.1859 | -101.929
6.1896 | -101.9026
6.1932 | -101.8705
6.1969 | -101.8411
6.2006 | -101.8068
6.2042 | -101.7661
6.2079 | -101.7227
6.2115 | -101.6768
6.2152 | -101.6303
6.2188 | -101.5883
6.2225 | -101.5514
6.2262 | -101.5246
6.2298 | -101.5062
6.2335 | -101.4976
6.2371 | -101.4975
6.2408 | -101.5021
6.2445 | -101.5176
6.2481 | -101.5421
6.2518 | -101.5723
6.2554 | -101.604
6.2591 | -101.6394
6.2627 | -101.6796
6.2664 | -101.716
6.2701 | -101.758
6.2737 | -101.7996
6.2774 | -101.8491
6.281 | -101.899
6.2847 | -101.9355
6.2884 | -101.9648
6.292 | -101.9812
6.2957 | -101.9969
6.2993 | -102.0019
6.303 | -101.9985
6.3066 | -101.99
6.3103 | -101.9749
6.314 | -101.9554
6.3176 | -101.9299
6.3213 | -101.9025
6.3249 | -101.8779
6.3286 | -101.8571
6.3323 | -101.8608
6.3359 | -101.8595
6.3396 | -101.8631
6.3432 | -101.8793
6.3469 | -101.903
6.3505 | -101.9427
6.3542 | -101.9924
6.3579 | -102.0457
6.3615 | -102.1023
6.3652 | -102.1595
6.3688 | -102.2179
6.3725 | -102.2614
6.3761 | -102.2989
6.3798 | -102.33
6.3835 | -102.3474
6.3871 | -102.3593
6.3908 | -102.3589
6.3944 | -102.3478
6.3981 | -102.3233
6.4018 | -102.2865
6.4054 | -102.243
6.4091 | -102.2034
6.4127 | -102.1608
6.4164 | -102.1209
6.42 | -102.0859
6.4237 | -102.0617
6.4274 | -102.0484
6.431 | -102.047
6.4347 | -102.0571
6.4383 | -102.0791
6.442 | -102.1041
6.4457 | -102.1367
6.4493 | -102.17
6.453 | -102.2042
6.4566 | -102.2407
6.4603 | -102.2743
6.4639 | -102.2959
6.4676 | -102.3126
6.4713 | -102.3271
6.4749 | -102.3275
6.4786 | -102.3274
6.4822 | -102.3204
6.4859 | -102.3065
6.4896 | -102.2875
6.4932 | -102.2718
6.4969 | -102.2549
6.5005 | -102.2439
6.5042 | -102.2367
6.5078 | -102.225
6.5115 | -102.2183
6.5152 | -102.2114
6.5188 | -102.2061
6.5225 | -102.2058
6.5261 | -102.2068
6.5298 | -102.2077
6.5334 | -102.2094
6.5371 | -102.2118
6.5408 | -102.2215
6.5444 | -102.2285
6.5481 | -102.2342
6.5517 | -102.2404
6.5554 | -102.243
6.5591 | -102.2449
6.5627 | -102.2493
6.5664 | -102.2475
6.57 | -102.2486
6.5737 | -102.2509
6.5773 | -102.252
6.581 | -102.2501
6.5847 | -102.2465
6.5883 | -102.2535
6.592 | -102.2604
6.5956 | -102.2696
6.5993 | -102.279
6.603 | -102.2908
6.6066 | -102.3113
6.6103 | -102.3309
6.6139 | -102.3458
6.6176 | -102.3641
6.6212 | -102.3803
6.6249 | -102.3933
6.6286 | -102.4073
6.6322 | -102.4219
6.6359 | -102.4335
6.6395 | -102.4338
6.6432 | -102.4395
6.6469 | -102.4419
6.6505 | -102.4372
6.6542 | -102.4221
6.6578 | -102.4016
6.6615 | -102.379
6.6651 | -102.3555
6.6688 | -102.3268
6.6725 | -102.2989
6.6761 | -102.2781
6.6798 | -102.2653
6.6834 | -102.2621
6.6871 | -102.2573
6.6907 | -102.2565
6.6944 | -102.2602
6.6981 | -102.2682
6.7017 | -102.2756
6.7054 | -102.2888
6.709 | -102.3039
6.7127 | -102.3201
6.7164 | -102.3293
6.72 | -102.3332
6.7237 | -102.3397
6.7273 | -102.3341
6.731 | -102.3277
6.7346 | -102.323
6.7383 | -102.3209
6.742 | -102.3183
6.7456 | -102.3174
6.7493 | -102.3235
6.7529 | -102.3378
6.7566 | -102.3489
6.7603 | -102.3581
6.7639 | -102.3677
6.7676 | -102.3776
6.7712 | -102.386
6.7749 | -102.382
6.7785 | -102.379
6.7822 | -102.3721
6.7859 | -102.3676
6.7895 | -102.3633
6.7932 | -102.354
6.7968 | -102.3419
6.8005 | -102.3223
6.8042 | -102.294
6.8078 | -102.2655
6.8115 | -102.2385
6.8151 | -102.2154
6.8188 | -102.1958
6.8224 | -102.182
6.8261 | -102.1803
6.8298 | -102.1816
6.8334 | -102.1929
6.8371 | -102.2157
6.8407 | -102.2429
6.8444 | -102.2779
6.848 | -102.3182
6.8517 | -102.3541
6.8554 | -102.3845
6.859 | -102.3808
6.8627 | -102.3894
6.8663 | -102.3882
6.87 | -102.3839
6.8737 | -102.3758
6.8773 | -102.3718
6.881 | -102.3683
6.8846 | -102.3639
6.8883 | -102.3663
6.8919 | -102.3835
6.8956 | -102.41
6.8993 | -102.4425
6.9029 | -102.4917
6.9066 | -102.5401
6.9102 | -102.589
6.9139 | -102.6393
6.9176 | -102.6853
6.9212 | -102.7297
6.9249 | -102.7614
6.9285 | -102.7805
6.9322 | -102.7912
6.9358 | -102.7959
6.9395 | -102.7932
6.9432 | -102.7863
6.9468 | -102.7714
6.9505 | -102.7462
6.9541 | -102.7224
6.9578 | -102.6959
6.9615 | -102.6659
6.9651 | -102.6459
6.9688 | -102.6289
6.9724 | -102.6116
6.9761 | -102.5974
6.9797 | -102.5908
6.9834 | -102.5879
6.9871 | -102.5923
6.9907 | -102.6089
6.9944 | -102.6293
6.998 | -102.654
7.0017 | -102.6874
7.0053 | -102.725
7.009 | -102.7598
7.0127 | -102.7975
7.0163 | -102.835
7.02 | -102.8678
7.0236 | -102.8978
7.0273 | -102.9205
7.031 | -102.9403
7.0346 | -102.956
7.0383 | -102.9628
7.0419 | -102.952
7.0456 | -102.9298
7.0492 | -102.8955
7.0529 | -102.8572
7.0566 | -102.813
7.0602 | -102.7635
7.0639 | -102.7164
7.0675 | -102.6759
7.0712 | -102.6343
7.0749 | -102.6013
7.0785 | -102.5801
7.0822 | -102.5671
7.0858 | -102.5675
7.0895 | -102.5736
7.0931 | -102.5814
7.0968 | -102.5935
7.1005 | -102.607
7.1041 | -102.6196
7.1078 | -102.6333
7.1114 | -102.6461
7.1151 | -102.6599
7.1188 | -102.6713
7.1224 | -102.6788
7.1261 | -102.6854
7.1297 | -102.6923
7.1334 | -102.7073
7.137 | -102.7308
7.1407 | -102.7513
7.1444 | -102.7743
7.148 | -102.8101
7.1517 | -102.8575
7.1553 | -102.9091
7.159 | -102.9606
7.1627 | -103.0118
7.1663 | -103.0559
7.17 | -103.0906
7.1736 | -103.116
7.1773 | -103.1297
7.1809 | -103.1338
7.1846 | -103.1255
7.1883 | -103.107
7.1919 | -103.0733
7.1956 | -103.0362
7.1992 | -102.9959
7.2029 | -102.9544
7.2065 | -102.9132
7.2102 | -102.874
7.2139 | -102.8381
7.2175 | -102.8107
7.2212 | -102.7928
7.2248 | -102.7792
7.2285 | -102.7711
7.2322 | -102.7633
7.2358 | -102.7582
7.2395 | -102.7533
7.2431 | -102.7487
7.2468 | -102.7521
7.2504 | -102.7534
7.2541 | -102.7511
7.2578 | -102.7492
7.2614 | -102.7535
7.2651 | -102.757
7.2687 | -102.7584
7.2724 | -102.7658
7.2761 | -102.7738
7.2797 | -102.7785
7.2834 | -102.7859
7.287 | -102.793
7.2907 | -102.8025
7.2943 | -102.8075
7.298 | -102.8139
7.3017 | -102.8203
7.3053 | -102.8261
7.309 | -102.8334
7.3126 | -102.8452
7.3163 | -102.8603
7.32 | -102.878
7.3236 | -102.8893
7.3273 | -102.8984
7.3309 | -102.9121
7.3346 | -102.9188
7.3382 | -102.927
7.3419 | -102.9278
7.3456 | -102.9232
7.3492 | -102.9147
7.3529 | -102.9019
7.3565 | -102.8841
7.3602 | -102.8609
7.3638 | -102.8398
7.3675 | -102.8195
7.3712 | -102.7972
7.3748 | -102.7766
7.3785 | -102.7474
7.3821 | -102.7222
7.3858 | -102.7133
7.3895 | -102.7038
7.3931 | -102.6969
7.3968 | -102.6888
7.4004 | -102.6885
7.4041 | -102.6975
7.4077 | -102.7109
7.4114 | -102.7283
7.4151 | -102.7516
7.4187 | -102.7769
7.4224 | -102.8047
7.426 | -102.8375
7.4297 | -102.8753
7.4334 | -102.9214
7.437 | -102.9709
7.4407 | -103.0212
7.4443 | -103.0753
7.448 | -103.1301
7.4516 | -103.1823
7.4553 | -103.231
7.459 | -103.2732
7.4626 | -103.3054
7.4663 | -103.3303
7.4699 | -103.3444
7.4736 | -103.3419
7.4773 | -103.3281
7.4809 | -103.3111
7.4846 | -103.2842
7.4882 | -103.2548
7.4919 | -103.2167
7.4955 | -103.177
7.4992 | -103.1345
7.5029 | -103.087
7.5065 | -103.0479
7.5102 | -103.0134
7.5138 | -102.9774
7.5175 | -102.9482
7.5211 | -102.9284
7.5248 | -102.9148
7.5285 | -102.9005
7.5321 | -102.8867
7.5358 | -102.88
7.5394 | -102.8727
7.5431 | -102.8661
7.5468 | -102.8567
7.5504 | -102.845
7.5541 | -102.8315
7.5577 | -102.8146
7.5614 | -102.7928
7.565 | -102.7693
7.5687 | -102.7432
7.5724 | -102.7178
7.576 | -102.695
7.5797 | -102.6724
7.5833 | -102.6585
7.587 | -102.6471
7.5907 | -102.6389
7.5943 | -102.6407
7.598 | -102.6535
7.6016 | -102.669
7.6053 | -102.6881
7.6089 | -102.7102
7.6126 | -102.7421
7.6163 | -102.7798
7.6199 | -102.8149
7.6236 | -102.8517
7.6272 | -102.8882
7.6309 | -102.9184
7.6346 | -102.9463
7.6382 | -102.9745
7.6419 | -103.0045
7.6455 | -103.0337
7.6492 | -103.053
7.6528 | -103.0671
7.6565 | -103.0765
7.6602 | -103.0875
7.6638 | -103.0986
7.6675 | -103.1066
7.6711 | -103.1153
7.6748 | -103.123
7.6784 | -103.1341
7.6821 | -103.1437
7.6858 | -103.1568
7.6894 | -103.1692
7.6931 | -103.177
7.6967 | -103.1798
7.7004 | -103.1899
7.7041 | -103.2005
7.7077 | -103.2047
7.7114 | -103.2106
7.715 | -103.2225
7.7187 | -103.2378
7.7223 | -103.2538
7.726 | -103.2685
7.7297 | -103.2893
7.7333 | -103.3198
7.737 | -103.3531
7.7406 | -103.3893
7.7443 | -103.4312
7.748 | -103.4774
7.7516 | -103.5217
7.7553 | -103.5633
7.7589 | -103.6023
7.7626 | -103.6433
7.7662 | -103.6816
7.7699 | -103.7004
7.7736 | -103.699
7.7772 | -103.6892
7.7809 | -103.67
7.7845 | -103.6351
7.7882 | -103.5918
7.7919 | -103.5421
7.7955 | -103.488
7.7992 | -103.4275
7.8028 | -103.3721
7.8065 | -103.3257
7.8101 | -103.2886
7.8138 | -103.2565
7.8175 | -103.2351
7.8211 | -103.2255
7.8248 | -103.2193
7.8284 | -103.2203
7.8321 | -103.2321
7.8357 | -103.2471
7.8394 | -103.2632
7.8431 | -103.2805
7.8467 | -103.2907
7.8504 | -103.3016
7.854 | -103.3123
7.8577 | -103.3162
7.8614 | -103.3167
7.865 | -103.3234
7.8687 | -103.331
7.8723 | -103.3356
7.876 | -103.3344
7.8796 | -103.336
7.8833 | -103.3468
7.887 | -103.3611
7.8906 | -103.3686
7.8943 | -103.377
7.8979 | -103.3912
7.9016 | -103.3997
7.9053 | -103.4083
7.9089 | -103.4165
7.9126 | -103.4186
7.9162 | -103.4031
7.9199 | -103.3979
7.9235 | -103.3859
7.9272 | -103.3726
7.9309 | -103.3544
7.9345 | -103.3359
7.9382 | -103.3185
7.9418 | -103.3013
7.9455 | -103.2822
7.9492 | -103.2606
7.9528 | -103.2434
7.9565 | -103.2263
7.9601 | -103.207
7.9638 | -103.1897
7.9674 | -103.17
7.9711 | -103.1487
7.9748 | -103.1289
7.9784 | -103.1081
7.9821 | -103.0835
7.9857 | -103.0621
7.9894 | -103.0462
7.9931 | -103.0321
7.9967 | -103.0174
8.0004 | -103.0061
8.004 | -103.0024
8.0077 | -103.0063
8.0113 | -103.0129
8.015 | -103.0251
8.0187 | -103.0459
8.0223 | -103.067
8.026 | -103.0944
8.0296 | -103.127
8.0333 | -103.1609
8.0369 | -103.1967
8.0406 | -103.2361
8.0443 | -103.2784
8.0479 | -103.3173
8.0516 | -103.3603
8.0552 | -103.4019
8.0589 | -103.4389
8.0626 | -103.477
8.0662 | -103.5127
8.0699 | -103.5469
8.0735 | -103.5781
8.0772 | -103.6103
8.0808 | -103.6425
8.0845 | -103.6679
8.0882 | -103.685
8.0918 | -103.7031
8.0955 | -103.718
8.0991 | -103.7285
8.1028 | -103.7365
8.1065 | -103.7383
8.1101 | -103.732
8.1138 | -103.7272
8.1174 | -103.7168
8.1211 | -103.6997
8.1247 | -103.6769
8.1284 | -103.6522
8.1321 | -103.6325
8.1357 | -103.6134
8.1394 | -103.593
8.143 | -103.5813
8.1467 | -103.5684
8.1503 | -103.5597
8.154 | -103.5529
8.1577 | -103.5453
8.1613 | -103.5387
8.165 | -103.533
8.1686 | -103.5281
8.1723 | -103.5178
8.176 | -103.5058
8.1796 | -103.4971
8.1833 | -103.4845
8.1869 | -103.4717
8.1906 | -103.4586
8.1942 | -103.4467
8.1979 | -103.4329
8.2016 | -103.4235
8.2052 | -103.4236
8.2089 | -103.4332
8.2125 | -103.4375
8.2162 | -103.4452
8.2199 | -103.4561
8.2235 | -103.475
8.2272 | -103.4985
8.2308 | -103.5265
8.2345 | -103.552
8.2381 | -103.5844
8.2418 | -103.619
8.2455 | -103.6512
8.2491 | -103.6827
8.2528 | -103.7077
8.2564 | -103.7334
8.2601 | -103.7534
8.2638 | -103.7652
8.2674 | -103.7767
8.2711 | -103.7794
8.2747 | -103.7722
8.2784 | -103.7585
8.282 | -103.7369
8.2857 | -103.7111
8.2894 | -103.6824
8.293 | -103.6518
8.2967 | -103.6184
8.3003 | -103.5798
8.304 | -103.5459
8.3077 | -103.5119
8.3113 | -103.4798
8.315 | -103.4521
8.3186 | -103.4337
8.3223 | -103.4192
8.3259 | -103.4079
8.3296 | -103.3993
8.3333 | -103.3998
8.3369 | -103.4043
8.3406 | -103.4122
8.3442 | -103.4186
8.3479 | -103.4273
8.3515 | -103.4408
8.3552 | -103.457
8.3589 | -103.4707
8.3625 | -103.4832
8.3662 | -103.4956
8.3698 | -103.5064
8.3735 | -103.5187
8.3772 | -103.5285
8.3808 | -103.5417
8.3845 | -103.5524
8.3881 | -103.5606
8.3918 | -103.5737
8.3954 | -103.5928
8.3991 | -103.6127
8.4028 | -103.6357
8.4064 | -103.6591
8.4101 | -103.6823
8.4137 | -103.7129
8.4174 | -103.7484
8.4211 | -103.7902
8.4247 | -103.8284
8.4284 | -103.8668
8.432 | -103.9015
8.4357 | -103.9314
8.4393 | -103.9535
8.443 | -103.9498
8.4467 | -103.9582
8.4503 | -103.9627
8.454 | -103.963
8.4576 | -103.9578
8.4613 | -103.9488
8.4649 | -103.9367
8.4686 | -103.9202
8.4723 | -103.9001
8.4759 | -103.8785
8.4796 | -103.8571
8.4832 | -103.8367
8.4869 | -103.8155
8.4906 | -103.8009
8.4942 | -103.7904
8.4979 | -103.7783
8.5015 | -103.7725
8.5052 | -103.7729
8.5088 | -103.7775
8.5125 | -103.7871
8.5162 | -103.8004
8.5198 | -103.8172
8.5235 | -103.8362
8.5271 | -103.853
8.5308 | -103.8757
8.5345 | -103.9023
8.5381 | -103.9258
8.5418 | -103.9425
8.5454 | -103.9578
8.5491 | -103.9763
8.5527 | -103.9946
8.5564 | -104.014
8.5601 | -104.0336
8.5637 | -104.0519
8.5674 | -104.0727
8.571 | -104.0934
8.5747 | -104.1125
8.5784 | -104.1318
8.582 | -104.1482
8.5857 | -104.1667
8.5893 | -104.188
8.593 | -104.2035
8.5966 | -104.2186
8.6003 | -104.2325
8.604 | -104.2374
8.6076 | -104.2387
8.6113 | -104.2356
8.6149 | -104.2295
8.6186 | -104.214
8.6223 | -104.1958
8.6259 | -104.1808
8.6296 | -104.154
8.6332 | -104.1266
8.6369 | -104.0988
8.6405 | -104.0734
8.6442 | -104.0481
8.6479 | -104.0224
8.6515 | -104.005
8.6552 | -103.9934
8.6588 | -103.9794
8.6625 | -103.966
8.6661 | -103.9571
8.6698 | -103.9529
8.6735 | -103.9445
8.6771 | -103.9425
8.6808 | -103.9395
8.6844 | -103.9389
8.6881 | -103.9408
8.6918 | -103.9489
8.6954 | -103.9571
8.6991 | -103.9627
8.7027 | -103.9714
8.7064 | -103.9837
8.71 | -104.0007
8.7137 | -104.0158
8.7174 | -104.0287
8.721 | -104.0437
8.7247 | -104.0605
8.7283 | -104.0756
8.732 | -104.0889
8.7357 | -104.1028
8.7393 | -104.117
8.743 | -104.134
8.7466 | -104.1463
8.7503 | -104.1552
8.7539 | -104.1613
8.7576 | -104.1728
8.7613 | -104.1795
8.7649 | -104.1875
8.7686 | -104.196
8.7722 | -104.2047
8.7759 | -104.2198
8.7796 | -104.2333
8.7832 | -104.2498
8.7869 | -104.2734
8.7905 | -104.3027
8.7942 | -104.3303
8.7978 | -104.3572
8.8015 | -104.3909
8.8052 | -104.4169
8.8088 | -104.4406
8.8125 | -104.4535
8.8161 | -104.4598
8.8198 | -104.4594
8.8234 | -104.4582
8.8271 | -104.4514
8.8308 | -104.4322
8.8344 | -104.4066
8.8381 | -104.3729
8.8417 | -104.334
8.8454 | -104.2933
8.8491 | -104.2533
8.8527 | -104.206
8.8564 | -104.1655
8.86 | -104.1317
8.8637 | -104.0984
8.8673 | -104.0671
8.871 | -104.0361
8.8747 | -104.0149
8.8783 | -104.0002
8.882 | -103.9902
8.8856 | -103.9796
8.8893 | -103.9761
8.893 | -103.9749
8.8966 | -103.9772
8.9003 | -103.9826
8.9039 | -103.9868
8.9076 | -103.9989
8.9112 | -104.0125
8.9149 | -104.025
8.9186 | -104.035
8.9222 | -104.0487
8.9259 | -104.0658
8.9295 | -104.076
8.9332 | -104.0859
8.9369 | -104.0935
8.9405 | -104.1053
8.9442 | -104.1152
8.9478 | -104.12
8.9515 | -104.1156
8.9551 | -104.1056
8.9588 | -104.097
8.9625 | -104.0825
8.9661 | -104.0692
8.9698 | -104.0445
8.9734 | -104.0266
8.9771 | -104.0055
8.9807 | -103.9823
8.9844 | -103.961
8.9881 | -103.9361
8.9917 | -103.9136
8.9954 | -103.8956
8.999 | -103.8809
9.0027 | -103.8759
9.0064 | -103.8725
9.01 | -103.8721
9.0137 | -103.8799
9.0173 | -103.8972
9.021 | -103.9176
9.0246 | -103.9428
9.0283 | -103.9689
9.032 | -104.0033
9.0356 | -104.0424
9.0393 | -104.0818
9.0429 | -104.1251
9.0466 | -104.1717
9.0503 | -104.2221
9.0539 | -104.2738
9.0576 | -104.3214
9.0612 | -104.3677
9.0649 | -104.4074
9.0685 | -104.4427
9.0722 | -104.4734
9.0759 | -104.4994
9.0795 | -104.5204
9.0832 | -104.5412
9.0868 | -104.5551
9.0905 | -104.5571
9.0942 | -104.5543
9.0978 | -104.5417
9.1015 | -104.5231
9.1051 | -104.5002
9.1088 | -104.4735
9.1124 | -104.4449
9.1161 | -104.4119
9.1198 | -104.3744
9.1234 | -104.3357
9.1271 | -104.2957
9.1307 | -104.2576
9.1344 | -104.2277
9.138 | -104.1979
9.1417 | -104.1745
9.1454 | -104.1601
9.149 | -104.1518
9.1527 | -104.148
9.1563 | -104.1522
9.16 | -104.1632
9.1637 | -104.181
9.1673 | -104.1998
9.171 | -104.2234
9.1746 | -104.2488
9.1783 | -104.2734
9.1819 | -104.2972
9.1856 | -104.3169
9.1893 | -104.3339
9.1929 | -104.3502
9.1966 | -104.3647
9.2002 | -104.3752
9.2039 | -104.3808
9.2076 | -104.3797
9.2112 | -104.3752
9.2149 | -104.3707
9.2185 | -104.3628
9.2222 | -104.3542
9.2258 | -104.3508
9.2295 | -104.343
9.2332 | -104.3391
9.2368 | -104.3394
9.2405 | -104.3422
9.2441 | -104.3454
9.2478 | -104.3505
9.2515 | -104.3612
9.2551 | -104.3761
9.2588 | -104.3888
9.2624 | -104.4005
9.2661 | -104.4142
9.2697 | -104.4243
9.2734 | -104.4358
9.2771 | -104.4465
9.2807 | -104.4534
9.2844 | -104.4574
9.288 | -104.463
9.2917 | -104.467
9.2953 | -104.4706
9.299 | -104.4731
9.3027 | -104.4727
9.3063 | -104.471
9.31 | -104.467
9.3136 | -104.4626
9.3173 | -104.4586
9.321 | -104.4525
9.3246 | -104.445
9.3283 | -104.4386
9.3319 | -104.4295
9.3356 | -104.4172
9.3392 | -104.4073
9.3429 | -104.3919
9.3466 | -104.3751
9.3502 | -104.357
9.3539 | -104.3382
9.3575 | -104.3208
9.3612 | -104.2985
9.3649 | -104.2809
9.3685 | -104.2601
9.3722 | -104.2407
9.3758 | -104.2236
9.3795 | -104.2119
9.3831 | -104.1972
9.3868 | -104.1847
9.3905 | -104.1727
9.3941 | -104.1662
9.3978 | -104.1613
9.4014 | -104.1565
9.4051 | -104.1539
9.4088 | -104.1495
9.4124 | -104.1455
9.4161 | -104.1431
9.4197 | -104.1409
9.4234 | -104.136
9.427 | -104.1287
9.4307 | -104.1271
9.4344 | -104.1276
9.438 | -104.1289
9.4417 | -104.1316
9.4453 | -104.1352
9.449 | -104.1416
9.4527 | -104.1476
9.4563 | -104.156
9.46 | -104.1657
9.4636 | -104.1707
9.4673 | -104.176
9.4709 | -104.1844
9.4746 | -104.1905
9.4783 | -104.1949
9.4819 | -104.1967
9.4856 | -104.2
9.4892 | -104.2039
9.4929 | -104.207
9.4965 | -104.2143
9.5002 | -104.2197
9.5039 | -104.2282
9.5075 | -104.2421
9.5112 | -104.2551
9.5148 | -104.2685
9.5185 | -104.2836
9.5222 | -104.3077
9.5258 | -104.3334
9.5295 | -104.361
9.5331 | -104.3857
9.5368 | -104.4113
9.5404 | -104.4375
9.5441 | -104.4596
9.5478 | -104.4785
9.5514 | -104.4979
9.5551 | -104.5158
9.5587 | -104.5305
9.5624 | -104.5434
9.5661 | -104.5558
9.5697 | -104.5601
9.5734 | -104.5625
9.577 | -104.568
9.5807 | -104.5718
9.5843 | -104.5775
9.588 | -104.5816
9.5917 | -104.5875
9.5953 | -104.5928
9.599 | -104.5997
9.6026 | -104.61
9.6063 | -104.6214
9.6099 | -104.6329
9.6136 | -104.6462
9.6173 | -104.6639
9.6209 | -104.6811
9.6246 | -104.6956
9.6282 | -104.7064
9.6319 | -104.7187
9.6356 | -104.7285
9.6392 | -104.741
9.6429 | -104.7524
9.6465 | -104.7592
9.6502 | -104.7641
9.6538 | -104.7711
9.6575 | -104.7802
9.6612 | -104.7861
9.6648 | -104.7868
9.6685 | -104.7862
9.6721 | -104.7852
9.6758 | -104.7852
9.6795 | -104.7863
9.6831 | -104.7874
9.6868 | -104.7891
9.6904 | -104.7939
9.6941 | -104.7955
9.6977 | -104.7962
9.7014 | -104.7954
9.7051 | -104.789
9.7087 | -104.7818
9.7124 | -104.7728
9.716 | -104.7642
9.7197 | -104.7563
9.7234 | -104.7482
9.727 | -104.7412
9.7307 | -104.7353
9.7343 | -104.7298
9.738 | -104.7253
9.7416 | -104.7212
9.7453 | -104.7204
9.749 | -104.7165
9.7526 | -104.7066
9.7563 | -104.6931
9.7599 | -104.6795
9.7636 | -104.6715
9.7673 | -104.6642
9.7709 | -104.6578
9.7746 | -104.6439
9.7782 | -104.6304
9.7819 | -104.6117
9.7855 | -104.5858
9.7892 | -104.5617
9.7929 | -104.5358
9.7965 | -104.5037
9.8002 | -104.4683
9.8038 | -104.4334
9.8075 | -104.395
9.8111 | -104.3583
9.8148 | -104.3227
9.8185 | -104.2876
9.8221 | -104.2515
9.8258 | -104.2157
9.8294 | -104.1888
9.8331 | -104.1667
9.8368 | -104.1489
9.8404 | -104.1385
9.8441 | -104.1365
9.8477 | -104.1449
9.8514 | -104.1576
9.855 | -104.1725
9.8587 | -104.19
9.8624 | -104.2097
9.866 | -104.2332
9.8697 | -104.2626
9.8733 | -104.2985
9.877 | -104.3376
9.8807 | -104.3735
9.8843 | -104.4095
9.888 | -104.4495
9.8916 | -104.4898
9.8953 | -104.5298
9.8989 | -104.5662
9.9026 | -104.596
9.9063 | -104.6179
9.9099 | -104.6337
9.9136 | -104.6457
9.9172 | -104.6582
9.9209 | -104.6629
9.9246 | -104.6632
9.9282 | -104.6643
9.9319 | -104.6542
9.9355 | -104.6431
9.9392 | -104.6306
9.9428 | -104.6159
9.9465 | -104.5994
9.9502 | -104.5833
9.9538 | -104.5661
9.9575 | -104.5511
9.9611 | -104.5372
9.9648 | -104.5247
9.9684 | -104.5155
9.9721 | -104.5071
9.9758 | -104.4994
9.9794 | -104.497
9.9831 | -104.4951
9.9867 | -104.495
9.9904 | -104.4966
9.9941 | -104.4996
9.9977 | -104.5084
10.0014 | -104.5162
10.005 | -104.5258
10.0087 | -104.5372
10.0123 | -104.5527
10.016 | -104.5689
10.0197 | -104.5873
10.0233 | -104.6116
10.027 | -104.6379
10.0306 | -104.6656
10.0343 | -104.6894
10.038 | -104.7143
10.0416 | -104.7397
10.0453 | -104.765
10.0489 | -104.7894
10.0526 | -104.8114
10.0562 | -104.8324
10.0599 | -104.8483
10.0636 | -104.8583
10.0672 | -104.8654
10.0709 | -104.8626
10.0745 | -104.8555
10.0782 | -104.849
10.0819 | -104.8367
10.0855 | -104.8205
10.0892 | -104.8011
10.0928 | -104.775
10.0965 | -104.7469
10.1001 | -104.7161
10.1038 | -104.6813
10.1075 | -104.643
10.1111 | -104.6037
10.1148 | -104.5653
10.1184 | -104.5245
10.1221 | -104.4835
10.1257 | -104.4421
10.1294 | -104.4036
10.1331 | -104.3715
10.1367 | -104.3417
10.1404 | -104.3146
10.144 | -104.2929
10.1477 | -104.2771
10.1514 | -104.2687
10.155 | -104.2648
10.1587 | -104.2599
10.1623 | -104.2632
10.166 | -104.2753
10.1696 | -104.2965
10.1733 | -104.3197
10.177 | -104.3441
10.1806 | -104.3696
10.1843 | -104.4003
10.1879 | -104.4334
10.1916 | -104.4641
10.1953 | -104.4942
10.1989 | -104.5245
10.2026 | -104.5534
10.2062 | -104.5803
10.2099 | -104.6044
10.2135 | -104.6251
10.2172 | -104.6412
10.2209 | -104.6537
10.2245 | -104.6623
10.2282 | -104.6641
10.2318 | -104.6592
10.2355 | -104.653
10.2392 | -104.6418
10.2428 | -104.6307
10.2465 | -104.6179
10.2501 | -104.6039
10.2538 | -104.593
10.2574 | -104.5844
10.2611 | -104.5759
10.2648 | -104.5666
10.2684 | -104.5613
10.2721 | -104.5606
10.2757 | -104.558
10.2794 | -104.5608
10.283 | -104.5649
10.2867 | -104.5704
10.2904 | -104.5776
10.294 | -104.5831
10.2977 | -104.5923
10.3013 | -104.5989
10.305 | -104.6062
10.3087 | -104.6173
10.3123 | -104.6289
10.316 | -104.6387
10.3196 | -104.6481
10.3233 | -104.6634
10.3269 | -104.6815
10.3306 | -104.6976
10.3343 | -104.7159
10.3379 | -104.7361
10.3416 | -104.7589
10.3452 | -104.7781
10.3489 | -104.7926
10.3526 | -104.8078
10.3562 | -104.8258
10.3599 | -104.8444
10.3635 | -104.8625
10.3672 | -104.8759
10.3708 | -104.8888
10.3745 | -104.8956
10.3782 | -104.9006
10.3818 | -104.906
10.3855 | -104.9092
10.3891 | -104.9109
10.3928 | -104.9109
10.3965 | -104.9078
10.4001 | -104.8975
10.4038 | -104.882
10.4074 | -104.8675
10.4111 | -104.8522
10.4147 | -104.8338
10.4184 | -104.8141
10.4221 | -104.7936
10.4257 | -104.7742
10.4294 | -104.7494
10.433 | -104.7242
10.4367 | -104.7037
10.4403 | -104.6843
10.444 | -104.664
10.4477 | -104.6448
10.4513 | -104.6309
10.455 | -104.6189
10.4586 | -104.6122
10.4623 | -104.6109
10.466 | -104.6119
10.4696 | -104.6146
10.4733 | -104.6197
10.4769 | -104.6294
10.4806 | -104.6439
10.4842 | -104.6577
10.4879 | -104.6742
10.4916 | -104.692
10.4952 | -104.7127
10.4989 | -104.7352
10.5025 | -104.7567
10.5062 | -104.7777
10.5099 | -104.7973
10.5135 | -104.8151
10.5172 | -104.8342
10.5208 | -104.8512
10.5245 | -104.8654
10.5281 | -104.8793
10.5318 | -104.8899
10.5355 | -104.8993
10.5391 | -104.907
10.5428 | -104.9114
10.5464 | -104.9138
10.5501 | -104.9188
10.5538 | -104.9162
10.5574 | -104.9211
10.5611 | -104.9232
10.5647 | -104.924
10.5684 | -104.9258
10.572 | -104.9259
10.5757 | -104.9232
10.5794 | -104.9273
10.583 | -104.9374
10.5867 | -104.9476
10.5903 | -104.9568
10.594 | -104.9707
10.5976 | -104.9883
10.6013 | -105.005
10.605 | -105.0228
10.6086 | -105.0424
10.6123 | -105.0579
10.6159 | -105.0738
10.6196 | -105.0897
10.6233 | -105.1091
10.6269 | -105.1254
10.6306 | -105.1354
10.6342 | -105.1409
10.6379 | -105.15
10.6415 | -105.1577
10.6452 | -105.1604
10.6489 | -105.1614
10.6525 | -105.1566
10.6562 | -105.1484
10.6598 | -105.1334
10.6635 | -105.1205
10.6672 | -105.1067
10.6708 | -105.09
10.6745 | -105.0702
10.6781 | -105.0457
10.6818 | -105.0167
10.6854 | -104.9854
10.6891 | -104.9548
10.6928 | -104.9252
10.6964 | -104.8926
10.7001 | -104.8617
10.7037 | -104.8298
10.7074 | -104.7993
10.7111 | -104.7703
10.7147 | -104.7414
10.7184 | -104.7121
10.722 | -104.6878
10.7257 | -104.664
10.7293 | -104.6436
10.733 | -104.6266
10.7367 | -104.6079
10.7403 | -104.5913
10.744 | -104.5781
10.7476 | -104.5686
10.7513 | -104.5638
10.7549 | -104.5634
10.7586 | -104.5647
10.7623 | -104.5643
10.7659 | -104.5666
10.7696 | -104.5716
10.7732 | -104.5769
10.7769 | -104.5871
10.7806 | -104.594
10.7842 | -104.6031
10.7879 | -104.6188
10.7915 | -104.637
10.7952 | -104.6494
10.7988 | -104.6637
10.8025 | -104.6824
10.8062 | -104.7004
10.8098 | -104.7207
10.8135 | -104.7427
10.8171 | -104.7663
10.8208 | -104.7844
10.8245 | -104.8013
10.8281 | -104.8205
10.8318 | -104.8399
10.8354 | -104.8628
10.8391 | -104.8816
10.8427 | -104.9001
10.8464 | -104.9144
10.8501 | -104.9286
10.8537 | -104.9472
10.8574 | -104.9638
10.861 | -104.9806
10.8647 | -104.9981
10.8684 | -105.0145
10.872 | -105.0294
10.8757 | -105.0473
10.8793 | -105.063
10.883 | -105.0792
10.8866 | -105.0927
10.8903 | -105.1036
10.894 | -105.113
10.8976 | -105.1179
10.9013 | -105.121
10.9049 | -105.1258
10.9086 | -105.1241
10.9123 | -105.12
10.9159 | -105.1194
10.9196 | -105.1174
10.9232 | -105.111
10.9269 | -105.1069
10.9305 | -105.101
10.9342 | -105.092
10.9379 | -105.0838
10.9415 | -105.0733
10.9452 | -105.0586
10.9488 | -105.0416
10.9525 | -105.0241
10.9561 | -105.0073
10.9598 | -104.9905
10.9635 | -104.9672
10.9671 | -104.9444
10.9708 | -104.9218
10.9744 | -104.9018
10.9781 | -104.8871
10.9818 | -104.8712
10.9854 | -104.8553
10.9891 | -104.8426
10.9927 | -104.829
10.9964 | -104.8212
11 | -104.8132
11.0037 | -104.8062
11.0074 | -104.7989
11.011 | -104.79
11.0147 | -104.7815
11.0183 | -104.7731
11.022 | -104.7664
11.0257 | -104.7629
11.0293 | -104.7621
11.033 | -104.7606
11.0366 | -104.758
11.0403 | -104.7581
11.0439 | -104.7585
11.0476 | -104.7604
11.0513 | -104.7656
11.0549 | -104.7682
11.0586 | -104.776
11.0622 | -104.7844
11.0659 | -104.7952
11.0696 | -104.8061
11.0732 | -104.8194
11.0769 | -104.8362
11.0805 | -104.8596
11.0842 | -104.8784
11.0878 | -104.8947
11.0915 | -104.9103
11.0952 | -104.9269
11.0988 | -104.9425
11.1025 | -104.9579
11.1061 | -104.972
11.1098 | -104.9832
11.1134 | -104.9955
11.1171 | -105.0055
11.1208 | -105.0121
11.1244 | -105.0204
11.1281 | -105.0221
11.1317 | -105.0172
11.1354 | -105.0093
11.1391 | -104.9981
11.1427 | -104.9861
11.1464 | -104.9746
11.15 | -104.9625
11.1537 | -104.9496
11.1573 | -104.9329
11.161 | -104.92
11.1647 | -104.9076
11.1683 | -104.8943
11.172 | -104.8824
11.1756 | -104.8714
11.1793 | -104.8616
11.183 | -104.8517
11.1866 | -104.8422
11.1903 | -104.8338
11.1939 | -104.8254
11.1976 | -104.8238
11.2012 | -104.823
11.2049 | -104.8281
11.2086 | -104.8371
11.2122 | -104.8445
11.2159 | -104.8573
11.2195 | -104.868
11.2232 | -104.8783
11.2269 | -104.892
11.2305 | -104.9063
11.2342 | -104.9172
11.2378 | -104.929
11.2415 | -104.9409
11.2451 | -104.9567
11.2488 | -104.9741
11.2525 | -104.9899
11.2561 | -105.0064
11.2598 | -105.0217
11.2634 | -105.037
11.2671 | -105.0507
11.2707 | -105.0708
11.2744 | -105.0891
11.2781 | -105.1008
11.2817 | -105.1064
11.2854 | -105.115
11.289 | -105.1273
11.2927 | -105.1408
11.2964 | -105.1526
11.3 | -105.1621
11.3037 | -105.1692
11.3073 | -105.1753
11.311 | -105.1765
11.3146 | -105.1751
11.3183 | -105.1741
11.322 | -105.1787
11.3256 | -105.1798
11.3293 | -105.1801
11.3329 | -105.1816
11.3366 | -105.1839
11.3403 | -105.1837
11.3439 | -105.1833
11.3476 | -105.1853
11.3512 | -105.1885
11.3549 | -105.1922
11.3585 | -105.1965
11.3622 | -105.1981
11.3659 | -105.2008
11.3695 | -105.1995
11.3732 | -105.198
11.3768 | -105.199
11.3805 | -105.1989
11.3842 | -105.1993
11.3878 | -105.2006
11.3915 | -105.1994
11.3951 | -105.1984
11.3988 | -105.1944
11.4024 | -105.1878
11.4061 | -105.1804
11.4098 | -105.1707
11.4134 | -105.1622
11.4171 | -105.155
11.4207 | -105.1447
11.4244 | -105.136
11.428 | -105.1274
11.4317 | -105.1177
11.4354 | -105.0985
11.439 | -105.0814
11.4427 | -105.0681
11.4463 | -105.0609
11.45 | -105.0479
11.4537 | -105.0304
11.4573 | -105.0141
11.461 | -104.9988
11.4646 | -104.9852
11.4683 | -104.9737
11.4719 | -104.9631
11.4756 | -104.9509
11.4793 | -104.9389
11.4829 | -104.9299
11.4866 | -104.9183
11.4902 | -104.91
11.4939 | -104.9047
11.4976 | -104.901
11.5012 | -104.8976
11.5049 | -104.8941
11.5085 | -104.8894
11.5122 | -104.8831
11.5158 | -104.878
11.5195 | -104.8726
11.5232 | -104.8686
11.5268 | -104.8647
11.5305 | -104.86
11.5341 | -104.854
11.5378 | -104.8448
11.5415 | -104.838
11.5451 | -104.8327
11.5488 | -104.8272
11.5524 | -104.8186
11.5561 | -104.8087
11.5597 | -104.8012
11.5634 | -104.7939
11.5671 | -104.7831
11.5707 | -104.7754
11.5744 | -104.767
11.578 | -104.7571
11.5817 | -104.745
11.5853 | -104.7366
11.589 | -104.7288
11.5927 | -104.7216
11.5963 | -104.7133
11.6 | -104.7092
11.6036 | -104.7096
11.6073 | -104.7106
11.611 | -104.7105
11.6146 | -104.7084
11.6183 | -104.7111
11.6219 | -104.7155
11.6256 | -104.7206
11.6292 | -104.7249
11.6329 | -104.7296
11.6366 | -104.7364
11.6402 | -104.7431
11.6439 | -104.7531
11.6475 | -104.7683
11.6512 | -104.7821
11.6549 | -104.7963
11.6585 | -104.812
11.6622 | -104.8284
11.6658 | -104.8461
11.6695 | -104.8642
11.6731 | -104.8884
11.6768 | -104.9146
11.6805 | -104.9436
11.6841 | -104.9717
11.6878 | -105.0015
11.6914 | -105.0322
11.6951 | -105.062
11.6988 | -105.0917
11.7024 | -105.1228
11.7061 | -105.1485
11.7097 | -105.1709
11.7134 | -105.1907
11.717 | -105.208
11.7207 | -105.2244
11.7244 | -105.2396
11.728 | -105.2469
11.7317 | -105.253
11.7353 | -105.2576
11.739 | -105.2514
11.7426 | -105.2414
11.7463 | -105.2291
11.75 | -105.2144
11.7536 | -105.1979
11.7573 | -105.1794
11.7609 | -105.1572
11.7646 | -105.1361
11.7683 | -105.1123
11.7719 | -105.0856
11.7756 | -105.0605
11.7792 | -105.0322
11.7829 | -105.0048
11.7865 | -104.981
11.7902 | -104.9576
11.7939 | -104.9383
11.7975 | -104.9202
11.8012 | -104.9039
11.8048 | -104.8926
11.8085 | -104.8801
11.8122 | -104.8725
11.8158 | -104.8662
11.8195 | -104.858
11.8231 | -104.8524
11.8268 | -104.8499
11.8304 | -104.8482
11.8341 | -104.8451
11.8378 | -104.8425
11.8414 | -104.8443
11.8451 | -104.8445
11.8487 | -104.844
11.8524 | -104.8422
11.8561 | -104.8413
11.8597 | -104.8411
11.8634 | -104.8385
11.867 | -104.8361
11.8707 | -104.8351
11.8743 | -104.8363
11.878 | -104.8421
11.8817 | -104.8475
11.8853 | -104.8525
11.889 | -104.8589
11.8926 | -104.8648
11.8963 | -104.8686
11.9 | -104.8744
11.9036 | -104.8834
11.9073 | -104.8937
11.9109 | -104.9062
11.9146 | -104.922
11.9182 | -104.9386
11.9219 | -104.9559
11.9256 | -104.9714
11.9292 | -104.9877
11.9329 | -105.0058
11.9365 | -105.0241
11.9402 | -105.0443
11.9438 | -105.0658
11.9475 | -105.0893
11.9512 | -105.1147
11.9548 | -105.1379
11.9585 | -105.1612
11.9621 | -105.1832
11.9658 | -105.205
11.9695 | -105.2264
11.9731 | -105.2476
11.9768 | -105.2631
11.9804 | -105.2826
11.9841 | -105.2967
11.9877 | -105.3106
11.9914 | -105.3219
11.9951 | -105.3319
11.9987 | -105.3361
12.0024 | -105.339
12.006 | -105.343
12.0097 | -105.3439
12.0134 | -105.3383
12.017 | -105.3301
12.0207 | -105.3229
12.0243 | -105.3143
12.028 | -105.307
12.0316 | -105.2992
12.0353 | -105.293
12.039 | -105.2839
12.0426 | -105.2747
12.0463 | -105.2666
12.0499 | -105.2569
12.0536 | -105.2479
12.0572 | -105.2447
12.0609 | -105.243
12.0646 | -105.2443
12.0682 | -105.2462
12.0719 | -105.2476
12.0755 | -105.2508
12.0792 | -105.2545
12.0829 | -105.2607
12.0865 | -105.2663
12.0902 | -105.2743
12.0938 | -105.2801
12.0975 | -105.284
12.1011 | -105.2851
12.1048 | -105.2856
12.1085 | -105.2852
12.1121 | -105.2832
12.1158 | -105.2784
12.1194 | -105.2756
12.1231 | -105.2726
12.1268 | -105.2651
12.1304 | -105.2595
12.1341 | -105.2547
12.1377 | -105.2491
12.1414 | -105.2467
12.145 | -105.2451
12.1487 | -105.2405
12.1524 | -105.2372
12.156 | -105.234
12.1597 | -105.2345
12.1633 | -105.2327
12.167 | -105.2317
12.1707 | -105.2319
12.1743 | -105.2337
12.178 | -105.2341
12.1816 | -105.2343
12.1853 | -105.2349
12.1889 | -105.2353
12.1926 | -105.2369
12.1963 | -105.2386
12.1999 | -105.2402
12.2036 | -105.2411
12.2072 | -105.2419
12.2109 | -105.2426
12.2145 | -105.2431
12.2182 | -105.2425
12.2219 | -105.2449
12.2255 | -105.2491
12.2292 | -105.2499
12.2328 | -105.2494
12.2365 | -105.2479
12.2402 | -105.2464
12.2438 | -105.2501
12.2475 | -105.2542
12.2511 | -105.2566
12.2548 | -105.2606
12.2584 | -105.2618
12.2621 | -105.2642
12.2658 | -105.2668
12.2694 | -105.2717
12.2731 | -105.2796
12.2767 | -105.2869
12.2804 | -105.2912
12.2841 | -105.2964
12.2877 | -105.2983
12.2914 | -105.2964
12.295 | -105.2912
12.2987 | -105.2889
12.3023 | -105.2849
12.306 | -105.2715
12.3097 | -105.2567
12.3133 | -105.2397
12.317 | -105.2208
12.3206 | -105.2011
12.3243 | -105.1764
12.328 | -105.1551
12.3316 | -105.1344
12.3353 | -105.1086
12.3389 | -105.0847
12.3426 | -105.0593
12.3462 | -105.0356
12.3499 | -105.0153
12.3536 | -104.996
12.3572 | -104.978
12.3609 | -104.9603
12.3645 | -104.9471
12.3682 | -104.936
12.3719 | -104.9255
12.3755 | -104.9187
12.3792 | -104.9143
12.3828 | -104.9126
12.3865 | -104.9151
12.3901 | -104.9213
12.3938 | -104.932
12.3975 | -104.9435
12.4011 | -104.9553
12.4048 | -104.9726
12.4084 | -104.9948
12.4121 | -105.0183
12.4157 | -105.0437
12.4194 | -105.07
12.4231 | -105.0971
12.4267 | -105.1221
12.4304 | -105.1477
12.434 | -105.1762
12.4377 | -105.2058
12.4414 | -105.2292
12.445 | -105.2503
12.4487 | -105.2696
12.4523 | -105.2842
12.456 | -105.2957
12.4596 | -105.3051
12.4633 | -105.3066
12.467 | -105.3045
12.4706 | -105.3021
12.4743 | -105.296
12.4779 | -105.2852
12.4816 | -105.2759
12.4853 | -105.2648
12.4889 | -105.2549
12.4926 | -105.243
12.4962 | -105.2281
12.4999 | -105.2089
12.5035 | -105.1927
12.5072 | -105.1782
12.5109 | -105.1591
12.5145 | -105.1453
12.5182 | -105.1313
12.5218 | -105.1234
12.5255 | -105.1162
12.5292 | -105.1077
12.5328 | -105.1009
12.5365 | -105.097
12.5401 | -105.0967
12.5438 | -105.0962
12.5474 | -105.0971
12.5511 | -105.1
12.5548 | -105.103
12.5584 | -105.109
12.5621 | -105.1155
12.5657 | -105.1195
12.5694 | -105.1231
12.573 | -105.1294
12.5767 | -105.1404
12.5804 | -105.1483
12.584 | -105.1543
12.5877 | -105.16
12.5913 | -105.1656
12.595 | -105.1724
12.5987 | -105.1829
12.6023 | -105.1887
12.606 | -105.1953
12.6096 | -105.1982
12.6133 | -105.1958
12.6169 | -105.1953
12.6206 | -105.2002
12.6243 | -105.2039
12.6279 | -105.2043
12.6316 | -105.2044
12.6352 | -105.2049
12.6389 | -105.2016
12.6426 | -105.199
12.6462 | -105.198
12.6499 | -105.1932
12.6535 | -105.1857
12.6572 | -105.1796
12.6608 | -105.1712
12.6645 | -105.161
12.6682 | -105.1565
12.6718 | -105.151
12.6755 | -105.147
12.6791 | -105.1443
12.6828 | -105.1415
12.6865 | -105.1394
12.6901 | -105.1366
12.6938 | -105.1376
12.6974 | -105.138
12.7011 | -105.1397
12.7047 | -105.1422
12.7084 | -105.1466
12.7121 | -105.1522
12.7157 | -105.1601
12.7194 | -105.1685
12.723 | -105.1774
12.7267 | -105.1857
12.7303 | -105.1933
12.734 | -105.2029
12.7377 | -105.212
12.7413 | -105.2232
12.745 | -105.2348
12.7486 | -105.2474
12.7523 | -105.2572
12.756 | -105.2643
12.7596 | -105.2705
12.7633 | -105.2775
12.7669 | -105.2862
12.7706 | -105.2926
12.7742 | -105.2969
12.7779 | -105.3009
12.7816 | -105.3
12.7852 | -105.2996
12.7889 | -105.3013
12.7925 | -105.3023
12.7962 | -105.304
12.7999 | -105.3038
12.8035 | -105.2999
12.8072 | -105.2919
12.8108 | -105.2811
12.8145 | -105.2707
12.8181 | -105.261
12.8218 | -105.2521
12.8255 | -105.2429
12.8291 | -105.2311
12.8328 | -105.2189
12.8364 | -105.2077
12.8401 | -105.1949
12.8438 | -105.1826
12.8474 | -105.17
12.8511 | -105.1555
12.8547 | -105.1397
12.8584 | -105.1256
12.862 | -105.1142
12.8657 | -105.1031
12.8694 | -105.0954
12.873 | -105.0867
12.8767 | -105.0752
12.8803 | -105.0663
12.884 | -105.0577
12.8876 | -105.0507
12.8913 | -105.0474
12.895 | -105.0454
12.8986 | -105.0455
12.9023 | -105.0484
12.9059 | -105.0532
12.9096 | -105.0585
12.9133 | -105.0655
12.9169 | -105.0761
12.9206 | -105.0853
12.9242 | -105.0949
12.9279 | -105.1047
12.9315 | -105.1162
12.9352 | -105.1261
12.9389 | -105.1359
12.9425 | -105.1448
12.9462 | -105.1502
12.9498 | -105.158
12.9535 | -105.168
12.9572 | -105.1805
12.9608 | -105.1921
12.9645 | -105.2023
12.9681 | -105.2116
12.9718 | -105.2216
12.9754 | -105.2296
12.9791 | -105.2409
12.9828 | -105.2516
12.9864 | -105.2611
12.9901 | -105.2706
12.9937 | -105.2869
12.9974 | -105.3053
13.0011 | -105.3172
13.0047 | -105.3236
13.0084 | -105.3326
13.012 | -105.341
13.0157 | -105.3496
13.0193 | -105.3606
13.023 | -105.3686
13.0267 | -105.3725
13.0303 | -105.3736
13.034 | -105.3792
13.0376 | -105.3847
13.0413 | -105.385
13.0449 | -105.3811
13.0486 | -105.3756
13.0523 | -105.3688
13.0559 | -105.3614
13.0596 | -105.3561
13.0632 | -105.3494
13.0669 | -105.339
13.0706 | -105.3304
13.0742 | -105.3206
13.0779 | -105.3064
13.0815 | -105.2926
13.0852 | -105.2805
13.0888 | -105.2699
13.0925 | -105.2599
13.0962 | -105.2491
13.0998 | -105.2429
13.1035 | -105.2362
13.1071 | -105.2297
13.1108 | -105.2242
13.1145 | -105.2213
13.1181 | -105.2199
13.1218 | -105.221
13.1254 | -105.2248
13.1291 | -105.2251
13.1327 | -105.2279
13.1364 | -105.2316
13.1401 | -105.2291
13.1437 | -105.2325
13.1474 | -105.2382
13.151 | -105.24
13.1547 | -105.2437
13.1584 | -105.2475
13.162 | -105.2471
13.1657 | -105.2455
13.1693 | -105.2437
13.173 | -105.2409
13.1766 | -105.2371
13.1803 | -105.2317
13.184 | -105.2278
13.1876 | -105.2242
13.1913 | -105.2131
13.1949 | -105.2076
13.1986 | -105.1982
13.2022 | -105.1908
13.2059 | -105.1792
13.2096 | -105.1657
13.2132 | -105.1563
13.2169 | -105.1513
13.2205 | -105.1478
13.2242 | -105.1392
13.2279 | -105.1305
13.2315 | -105.126
13.2352 | -105.1225
13.2388 | -105.1188
13.2425 | -105.1177
13.2461 | -105.1173
13.2498 | -105.1142
13.2535 | -105.1142
13.2571 | -105.1161
13.2608 | -105.1167
13.2644 | -105.121
13.2681 | -105.1275
13.2718 | -105.1335
13.2754 | -105.139
13.2791 | -105.1476
13.2827 | -105.1597
13.2864 | -105.1765
13.29 | -105.1928
13.2937 | -105.2111
13.2974 | -105.23
13.301 | -105.248
13.3047 | -105.2651
13.3083 | -105.2834
13.312 | -105.3021
13.3157 | -105.3232
13.3193 | -105.3456
13.323 | -105.3656
13.3266 | -105.3817
13.3303 | -105.3988
13.3339 | -105.4095
13.3376 | -105.4214
13.3413 | -105.4311
13.3449 | -105.4403
13.3486 | -105.4485
13.3522 | -105.4577
13.3559 | -105.4608
13.3596 | -105.4606
13.3632 | -105.4579
13.3669 | -105.456
13.3705 | -105.4496
13.3742 | -105.441
13.3778 | -105.4311
13.3815 | -105.4195
13.3852 | -105.4056
13.3888 | -105.3926
13.3925 | -105.3801
13.3961 | -105.3633
13.3998 | -105.3477
13.4034 | -105.3304
13.4071 | -105.3108
13.4108 | -105.2898
13.4144 | -105.2711
13.4181 | -105.254
13.4217 | -105.2393
13.4254 | -105.2249
13.4291 | -105.2107
13.4327 | -105.1978
13.4364 | -105.1828
13.44 | -105.1645
13.4437 | -105.1536
13.4473 | -105.1443
13.451 | -105.1347
13.4547 | -105.1292
13.4583 | -105.1227
13.462 | -105.1166
13.4656 | -105.1153
13.4693 | -105.1137
13.473 | -105.1085
13.4766 | -105.105
13.4803 | -105.098
13.4839 | -105.0901
13.4876 | -105.0878
13.4912 | -105.0864
13.4949 | -105.0868
13.4986 | -105.0868
13.5022 | -105.0871
13.5059 | -105.0879
13.5095 | -105.0861
13.5132 | -105.0857
13.5168 | -105.0863
13.5205 | -105.0863
13.5242 | -105.0856
13.5278 | -105.086
13.5315 | -105.085
13.5351 | -105.0868
13.5388 | -105.0882
13.5425 | -105.0888
13.5461 | -105.0891
13.5498 | -105.0915
13.5534 | -105.0965
13.5571 | -105.101
13.5607 | -105.1025
13.5644 | -105.107
13.5681 | -105.1128
13.5717 | -105.1189
13.5754 | -105.1279
13.579 | -105.1373
13.5827 | -105.1418
13.5864 | -105.1468
13.59 | -105.1531
13.5937 | -105.158
13.5973 | -105.1639
13.601 | -105.1692
13.6046 | -105.1724
13.6083 | -105.1762
13.612 | -105.1813
13.6156 | -105.1832
13.6193 | -105.184
13.6229 | -105.1837
13.6266 | -105.1813
13.6303 | -105.1796
13.6339 | -105.1823
13.6376 | -105.1741
13.6412 | -105.1625
13.6449 | -105.1542
13.6485 | -105.145
13.6522 | -105.1352
13.6559 | -105.1273
13.6595 | -105.1165
13.6632 | -105.1075
13.6668 | -105.0997
13.6705 | -105.0922
13.6742 | -105.0871
13.6778 | -105.0822
13.6815 | -105.0749
13.6851 | -105.071
13.6888 | -105.0654
13.6924 | -105.0612
13.6961 | -105.0577
13.6998 | -105.0557
13.7034 | -105.0539
13.7071 | -105.0543
13.7107 | -105.0573
13.7144 | -105.0631
13.718 | -105.0781
13.7217 | -105.0856
13.7254 | -105.0963
13.729 | -105.1112
13.7327 | -105.1266
13.7363 | -105.1412
13.74 | -105.1565
13.7437 | -105.1728
13.7473 | -105.1873
13.751 | -105.2038
13.7546 | -105.2247
13.7583 | -105.2481
13.7619 | -105.2695
13.7656 | -105.2913
13.7693 | -105.3139
13.7729 | -105.3357
13.7766 | -105.3625
13.7802 | -105.3906
13.7839 | -105.4162
13.7876 | -105.4401
13.7912 | -105.4656
13.7949 | -105.4886
13.7985 | -105.5105
13.8022 | -105.5308
13.8058 | -105.5506
13.8095 | -105.5678
13.8132 | -105.5823
13.8168 | -105.5972
13.8205 | -105.6135
13.8241 | -105.6245
13.8278 | -105.6303
13.8315 | -105.6268
13.8351 | -105.6279
13.8388 | -105.6271
13.8424 | -105.6186
13.8461 | -105.6046
13.8497 | -105.5896
13.8534 | -105.5732
13.8571 | -105.5555
13.8607 | -105.5334
13.8644 | -105.5088
13.868 | -105.4797
13.8717 | -105.4501
13.8753 | -105.4192
13.879 | -105.3859
13.8827 | -105.3548
13.8863 | -105.3276
13.89 | -105.3001
13.8936 | -105.2735
13.8973 | -105.2498
13.901 | -105.226
13.9046 | -105.206
13.9083 | -105.1865
13.9119 | -105.1694
13.9156 | -105.1525
13.9192 | -105.1434
13.9229 | -105.141
13.9266 | -105.1378
13.9302 | -105.1381
13.9339 | -105.1413
13.9375 | -105.1464
13.9412 | -105.1533
13.9449 | -105.1664
13.9485 | -105.1843
13.9522 | -105.2008
13.9558 | -105.2155
13.9595 | -105.219
13.9631 | -105.2078
13.9668 | -105.2026
13.9705 | -105.2044
13.9741 | -105.207
13.9778 | -105.2037
13.9814 | -105.1985
13.9851 | -105.1909
13.9888 | -105.1795
13.9924 | -105.1658
13.9961 | -105.1468
13.9997 | -105.1199
14.0034 | -105.092
14.007 | -105.0579
14.0107 | -105.0212
14.0144 | -104.9836
14.018 | -104.9441
14.0217 | -104.8999
14.0253 | -104.8563
14.029 | -104.8104
14.0326 | -104.7669
14.0363 | -104.7229
14.04 | -104.6793
14.0436 | -104.64
14.0473 | -104.6074
14.0509 | -104.5775
14.0546 | -104.5503
14.0583 | -104.5272
14.0619 | -104.5109
14.0656 | -104.5008
14.0692 | -104.496
14.0729 | -104.4971
14.0765 | -104.5036
14.0802 | -104.515
14.0839 | -104.5293
14.0875 | -104.5481
14.0912 | -104.569
14.0948 | -104.5934
14.0985 | -104.6187
14.1022 | -104.6459
14.1058 | -104.6699
14.1095 | -104.6976
14.1131 | -104.7282
14.1168 | -104.7583
14.1204 | -104.7855
14.1241 | -104.8108
14.1278 | -104.8361
14.1314 | -104.859
14.1351 | -104.8826
14.1387 | -104.9069
14.1424 | -104.9314
14.1461 | -104.9532
14.1497 | -104.9752
14.1534 | -104.9959
14.157 | -105.0186
14.1607 | -105.0496
14.1643 | -105.0913
14.168 | -105.1211
14.1717 | -105.1408
14.1753 | -105.1565
14.179 | -105.1736
14.1826 | -105.1944
14.1863 | -105.2166
14.1899 | -105.2392
14.1936 | -105.2601
14.1973 | -105.2827
14.2009 | -105.3067
14.2046 | -105.328
14.2082 | -105.3496
14.2119 | -105.3692
14.2156 | -105.3868
14.2192 | -105.4021
14.2229 | -105.4172
14.2265 | -105.4321
14.2302 | -105.4484
14.2338 | -105.4602
14.2375 | -105.4686
14.2412 | -105.478
14.2448 | -105.4831
14.2485 | -105.4771
14.2521 | -105.4788
14.2558 | -105.4808
14.2595 | -105.4806
14.2631 | -105.4793
14.2668 | -105.4784
14.2704 | -105.4767
14.2741 | -105.4749
14.2777 | -105.4702
14.2814 | -105.4649
14.2851 | -105.4597
14.2887 | -105.4554
14.2924 | -105.4509
14.296 | -105.4512
14.2997 | -105.4499
14.3034 | -105.4514
14.307 | -105.454
14.3107 | -105.4562
14.3143 | -105.4537
14.318 | -105.4513
14.3216 | -105.4521
14.3253 | -105.4528
14.329 | -105.4535
14.3326 | -105.453
14.3363 | -105.4518
14.3399 | -105.4493
14.3436 | -105.4481
14.3473 | -105.4489
14.3509 | -105.4487
14.3546 | -105.445
14.3582 | -105.446
14.3619 | -105.4477
14.3655 | -105.4505
14.3692 | -105.449
14.3729 | -105.4471
14.3765 | -105.4466
14.3802 | -105.4465
14.3838 | -105.4483
14.3875 | -105.4488
14.3911 | -105.4465
14.3948 | -105.4446
14.3985 | -105.4438
14.4021 | -105.4405
14.4058 | -105.4391
14.4094 | -105.4381
14.4131 | -105.4372
14.4168 | -105.4352
14.4204 | -105.4348
14.4241 | -105.4334
14.4277 | -105.4312
14.4314 | -105.4302
14.435 | -105.4299
14.4387 | -105.4272
14.4424 | -105.4224
14.446 | -105.4174
14.4497 | -105.4146
14.4533 | -105.4163
14.457 | -105.4143
14.4607 | -105.4081
14.4643 | -105.4028
14.468 | -105.4018
14.4716 | -105.4017
14.4753 | -105.3956
14.4789 | -105.3889
14.4826 | -105.3815
14.4863 | -105.3774
14.4899 | -105.3702
14.4936 | -105.3612
14.4972 | -105.3506
14.5009 | -105.3417
14.5046 | -105.3313
14.5082 | -105.3227
14.5119 | -105.3091
14.5155 | -105.2964
14.5192 | -105.2851
14.5228 | -105.274
14.5265 | -105.2623
14.5302 | -105.2507
14.5338 | -105.2397
14.5375 | -105.2308
14.5411 | -105.2207
14.5448 | -105.2134
14.5484 | -105.2101
14.5521 | -105.2053
14.5558 | -105.1988
14.5594 | -105.1936
14.5631 | -105.1909
14.5667 | -105.1859
14.5704 | -105.1836
14.5741 | -105.1822
14.5777 | -105.1861
14.5814 | -105.1902
14.585 | -105.1916
14.5887 | -105.1956
14.5923 | -105.1987
14.596 | -105.2037
14.5997 | -105.2115
14.6033 | -105.2192
14.607 | -105.2253
14.6106 | -105.2297
14.6143 | -105.2295
14.618 | -105.2303
14.6216 | -105.229
14.6253 | -105.2283
14.6289 | -105.2265
14.6326 | -105.2238
14.6362 | -105.22
14.6399 | -105.2156
14.6436 | -105.2075
14.6472 | -105.2035
14.6509 | -105.2
14.6545 | -105.1951
14.6582 | -105.1894
14.6618 | -105.1806
14.6655 | -105.1705
14.6692 | -105.1615
14.6728 | -105.1538
14.6765 | -105.1486
14.6801 | -105.1416
14.6838 | -105.1389
14.6875 | -105.1371
14.6911 | -105.1368
14.6948 | -105.135
14.6984 | -105.1362
14.7021 | -105.1364
14.7057 | -105.1335
14.7094 | -105.1316
14.7131 | -105.1318
14.7167 | -105.1328
14.7204 | -105.136
14.724 | -105.1383
14.7277 | -105.1395
14.7314 | -105.1367
14.735 | -105.1322
14.7387 | -105.1279
14.7423 | -105.1228
14.746 | -105.1175
14.7496 | -105.1113
14.7533 | -105.1008
14.757 | -105.088
14.7606 | -105.0768
14.7643 | -105.067
14.7679 | -105.0555
14.7716 | -105.0455
14.7753 | -105.0294
14.7789 | -105.018
14.7826 | -105.0033
14.7862 | -104.9906
14.7899 | -104.9781
14.7935 | -104.9643
14.7972 | -104.9536
14.8009 | -104.9431
14.8045 | -104.9339
14.8082 | -104.9222
14.8118 | -104.9101
14.8155 | -104.9013
14.8192 | -104.8959
14.8228 | -104.8947
14.8265 | -104.8933
14.8301 | -104.8892
14.8338 | -104.8879
14.8374 | -104.8882
14.8411 | -104.8901
14.8448 | -104.892
14.8484 | -104.8956
14.8521 | -104.8955
14.8557 | -104.8979
14.8594 | -104.9009
14.863 | -104.9034
14.8667 | -104.9014
14.8704 | -104.8973
14.874 | -104.8967
14.8777 | -104.9011
14.8813 | -104.9047
14.885 | -104.9058
14.8887 | -104.9036
14.8923 | -104.9021
14.896 | -104.8983
14.8996 | -104.8936
14.9033 | -104.8886
14.9069 | -104.8831
14.9106 | -104.8776
14.9143 | -104.8713
14.9179 | -104.8645
14.9216 | -104.855
14.9252 | -104.8433
14.9289 | -104.8321
14.9326 | -104.8224
14.9362 | -104.8123
14.9399 | -104.8032
14.9435 | -104.7964
14.9472 | -104.7903
14.9508 | -104.7843
14.9545 | -104.7811
14.9582 | -104.7782
14.9618 | -104.7771
14.9655 | -104.7796
14.9691 | -104.7832
14.9728 | -104.7859
14.9765 | -104.7905
14.9801 | -104.7964
14.9838 | -104.8003
14.9874 | -104.8047
14.9911 | -104.8101
14.9947 | -104.8167
14.9984 | -104.8234
15.0021 | -104.8321
15.0057 | -104.8394
15.0094 | -104.8482
15.013 | -104.8561
15.0167 | -104.8599
15.0203 | -104.8626
15.024 | -104.8646
15.0277 | -104.866
15.0313 | -104.8634
15.035 | -104.8593
15.0386 | -104.8526
15.0423 | -104.8461
15.046 | -104.8362
15.0496 | -104.8225
15.0533 | -104.8096
15.0569 | -104.7924
15.0606 | -104.7749
15.0642 | -104.7547
15.0679 | -104.738
15.0716 | -104.7203
15.0752 | -104.7021
15.0789 | -104.6857
15.0825 | -104.6725
15.0862 | -104.6597
15.0899 | -104.6436
15.0935 | -104.6269
15.0972 | -104.6117
15.1008 | -104.5979
15.1045 | -104.5835
15.1081 | -104.5727
15.1118 | -104.5655
15.1155 | -104.5597
15.1191 | -104.5559
15.1228 | -104.5528
15.1264 | -104.5535
15.1301 | -104.5581
15.1338 | -104.567
15.1374 | -104.5753
15.1411 | -104.5844
15.1447 | -104.5939
15.1484 | -104.6042
15.152 | -104.6143
15.1557 | -104.624
15.1594 | -104.6359
15.163 | -104.6494
15.1667 | -104.663
15.1703 | -104.6743
15.174 | -104.6871
15.1776 | -104.7016
15.1813 | -104.7172
15.185 | -104.7316
15.1886 | -104.7452
15.1923 | -104.7578
15.1959 | -104.7715
15.1996 | -104.7871
15.2033 | -104.8034
15.2069 | -104.8171
15.2106 | -104.8272
15.2142 | -104.8371
15.2179 | -104.845
15.2215 | -104.8536
15.2252 | -104.8612
15.2289 | -104.868
15.2325 | -104.8759
15.2362 | -104.8828
15.2398 | -104.8896
15.2435 | -104.8982
15.2472 | -104.9072
15.2508 | -104.9133
15.2545 | -104.9201
15.2581 | -104.926
15.2618 | -104.9337
15.2654 | -104.9422
15.2691 | -104.9497
15.2728 | -104.9588
15.2764 | -104.9674
15.2801 | -104.9783
15.2837 | -104.9874
15.2874 | -104.9951
15.2911 | -104.9983
15.2947 | -105.0005
15.2984 | -105.0029
15.302 | -105.0081
15.3057 | -105.0127
15.3093 | -105.018
15.313 | -105.0185
15.3167 | -105.0197
15.3203 | -105.0229
15.324 | -105.025
15.3276 | -105.0225
15.3313 | -105.0213
15.3349 | -105.0221
15.3386 | -105.0225
15.3423 | -105.0208
15.3459 | -105.0174
15.3496 | -105.0127
15.3532 | -105.0078
15.3569 | -105.0038
15.3606 | -105.001
15.3642 | -104.9971
15.3679 | -104.9934
15.3715 | -104.9932
15.3752 | -104.9949
15.3788 | -104.9939
15.3825 | -104.9911
15.3862 | -104.9883
15.3898 | -104.9878
15.3935 | -104.9855
15.3971 | -104.9821
15.4008 | -104.9772
15.4045 | -104.9747
15.4081 | -104.9705
15.4118 | -104.9674
15.4154 | -104.966
15.4191 | -104.962
15.4227 | -104.9581
15.4264 | -104.9544
15.4301 | -104.9512
15.4337 | -104.9477
15.4374 | -104.9411
15.441 | -104.9342
15.4447 | -104.9294
15.4484 | -104.9216
15.452 | -104.9155
15.4557 | -104.9121
15.4593 | -104.9081
15.463 | -104.9018
15.4666 | -104.8957
15.4703 | -104.8891
15.474 | -104.8844
15.4776 | -104.8808
15.4813 | -104.8761
15.4849 | -104.8731
15.4886 | -104.8661
15.4922 | -104.8597
15.4959 | -104.8572
15.4996 | -104.8532
15.5032 | -104.8494
15.5069 | -104.8465
15.5105 | -104.8457
15.5142 | -104.8448
15.5179 | -104.8452
15.5215 | -104.8448
15.5252 | -104.8443
15.5288 | -104.8448
15.5325 | -104.8444
15.5361 | -104.8461
15.5398 | -104.8464
15.5435 | -104.8454
15.5471 | -104.847
15.5508 | -104.8517
15.5544 | -104.853
15.5581 | -104.8541
15.5618 | -104.855
15.5654 | -104.8569
15.5691 | -104.8564
15.5727 | -104.8552
15.5764 | -104.8534
15.58 | -104.8476
15.5837 | -104.8372
15.5874 | -104.8275
15.591 | -104.8164
15.5947 | -104.8027
15.5983 | -104.7849
15.602 | -104.7693
15.6057 | -104.7522
15.6093 | -104.7334
15.613 | -104.7151
15.6166 | -104.6976
15.6203 | -104.6786
15.6239 | -104.6575
15.6276 | -104.631
15.6313 | -104.6067
15.6349 | -104.5831
15.6386 | -104.5601
15.6422 | -104.5376
15.6459 | -104.5126
15.6495 | -104.4887
15.6532 | -104.4672
15.6569 | -104.4445
15.6605 | -104.4219
15.6642 | -104.4019
15.6678 | -104.3825
15.6715 | -104.3649
15.6752 | -104.3471
15.6788 | -104.3259
15.6825 | -104.3064
15.6861 | -104.2897
15.6898 | -104.2725
15.6934 | -104.2566
15.6971 | -104.2404
15.7008 | -104.2236
15.7044 | -104.2081
15.7081 | -104.1906
15.7117 | -104.178
15.7154 | -104.1651
15.7191 | -104.1517
15.7227 | -104.1413
15.7264 | -104.1312
15.73 | -104.1216
15.7337 | -104.1138
15.7373 | -104.1083
15.741 | -104.1028
15.7447 | -104.0952
15.7483 | -104.0884
15.752 | -104.0852
15.7556 | -104.083
15.7593 | -104.083
15.763 | -104.0834
15.7666 | -104.0843
15.7703 | -104.0868
15.7739 | -104.0909
15.7776 | -104.0945
15.7812 | -104.0972
15.7849 | -104.1016
15.7886 | -104.1082
15.7922 | -104.1159
15.7959 | -104.1222
15.7995 | -104.129
15.8032 | -104.1392
15.8069 | -104.151
15.8105 | -104.1618
15.8142 | -104.1679
15.8178 | -104.1737
15.8215 | -104.1827
15.8251 | -104.1906
15.8288 | -104.1969
15.8325 | -104.2034
15.8361 | -104.2106
15.8398 | -104.2149
15.8434 | -104.2184
15.8471 | -104.2219
15.8507 | -104.2274
15.8544 | -104.2357
15.8581 | -104.2418
15.8617 | -104.2459
15.8654 | -104.2482
15.869 | -104.2501
15.8727 | -104.253
15.8764 | -104.2521
15.88 | -104.25
15.8837 | -104.2501
15.8873 | -104.2512
15.891 | -104.248
15.8946 | -104.2417
15.8983 | -104.2337
15.902 | -104.224
15.9056 | -104.2155
15.9093 | -104.2029
15.9129 | -104.1871
15.9166 | -104.1717
15.9203 | -104.1556
15.9239 | -104.136
15.9276 | -104.1141
15.9312 | -104.0896
15.9349 | -104.0648
15.9385 | -104.0389
15.9422 | -104.0117
15.9459 | -103.9805
15.9495 | -103.9477
15.9532 | -103.9151
15.9568 | -103.8828
15.9605 | -103.847
15.9642 | -103.8105
15.9678 | -103.7729
15.9715 | -103.7351
15.9751 | -103.6975
15.9788 | -103.6582
15.9824 | -103.6196
15.9861 | -103.5792
15.9898 | -103.5412
15.9934 | -103.5057
15.9971 | -103.4693
16.0007 | -103.4339
16.0044 | -103.4006
16.008 | -103.3691
16.0117 | -103.3404
16.0154 | -103.3141
16.019 | -103.2904
16.0227 | -103.2704
16.0263 | -103.2499
16.03 | -103.2324
16.0337 | -103.2172
16.0373 | -103.2003
16.041 | -103.1876
16.0446 | -103.1775
16.0483 | -103.1692
16.0519 | -103.1644
16.0556 | -103.1616
16.0593 | -103.1594
16.0629 | -103.1576
16.0666 | -103.1588
16.0702 | -103.1613
16.0739 | -103.1651
16.0776 | -103.1716
16.0812 | -103.1778
16.0849 | -103.1852
16.0885 | -103.1953
16.0922 | -103.206
16.0958 | -103.2168
16.0995 | -103.2303
16.1032 | -103.2453
16.1068 | -103.2621
16.1105 | -103.28
16.1141 | -103.2982
16.1178 | -103.3174
16.1214 | -103.3388
16.1251 | -103.3624
16.1288 | -103.3877
16.1324 | -103.4119
16.1361 | -103.4379
16.1397 | -103.4684
16.1434 | -103.5002
16.1471 | -103.5323
16.1507 | -103.5642
16.1544 | -103.5978
16.158 | -103.6337
16.1617 | -103.6694
16.1653 | -103.7038
16.169 | -103.7353
16.1727 | -103.7658
16.1763 | -103.7986
16.18 | -103.8309
16.1836 | -103.8619
16.1873 | -103.8919
16.191 | -103.9218
16.1946 | -103.9497
16.1983 | -103.9771
16.2019 | -104.0031
16.2056 | -104.0273
16.2092 | -104.0496
16.2129 | -104.0687
16.2166 | -104.0862
16.2202 | -104.1011
16.2239 | -104.1133
16.2275 | -104.1258
16.2312 | -104.1365
16.2349 | -104.1488
16.2385 | -104.1597
16.2422 | -104.1696
16.2458 | -104.1796
16.2495 | -104.1886
16.2531 | -104.1996
16.2568 | -104.2112
16.2605 | -104.221
16.2641 | -104.2314
16.2678 | -104.2461
16.2714 | -104.2611
16.2751 | -104.2753
16.2788 | -104.2884
16.2824 | -104.3
16.2861 | -104.3132
16.2897 | -104.3309
16.2934 | -104.3496
16.297 | -104.3689
16.3007 | -104.3875
16.3044 | -104.4048
16.308 | -104.4239
16.3117 | -104.4442
16.3153 | -104.4643
16.319 | -104.4826
16.3226 | -104.5034
16.3263 | -104.5234
16.33 | -104.5404
16.3336 | -104.5586
16.3373 | -104.5752
16.3409 | -104.5907
16.3446 | -104.6044
16.3483 | -104.615
16.3519 | -104.6256
16.3556 | -104.6371
16.3592 | -104.634
16.3629 | -104.6426
16.3665 | -104.6496
16.3702 | -104.6504
16.3739 | -104.6471
16.3775 | -104.6415
16.3812 | -104.6352
16.3848 | -104.629
16.3885 | -104.621
16.3922 | -104.6126
16.3958 | -104.6043
16.3995 | -104.5951
16.4031 | -104.5877
16.4068 | -104.5804
16.4104 | -104.5702
16.4141 | -104.5599
16.4178 | -104.549
16.4214 | -104.5403
16.4251 | -104.535
16.4287 | -104.5281
16.4324 | -104.5219
16.4361 | -104.5184
16.4397 | -104.5134
16.4434 | -104.5107
16.447 | -104.5104
16.4507 | -104.5133
16.4543 | -104.5151
16.458 | -104.5179
16.4617 | -104.5221
16.4653 | -104.5286
16.469 | -104.5343
16.4726 | -104.5439
16.4763 | -104.5534
16.4799 | -104.5632
16.4836 | -104.5759
16.4873 | -104.5888
16.4909 | -104.6068
16.4946 | -104.627
16.4982 | -104.6454
16.5019 | -104.6643
16.5056 | -104.686
16.5092 | -104.7086
16.5129 | -104.7295
16.5165 | -104.751
16.5202 | -104.7693
16.5238 | -104.7828
16.5275 | -104.7963
16.5312 | -104.8095
16.5348 | -104.825
16.5385 | -104.8379
16.5421 | -104.8426
16.5458 | -104.8464
16.5495 | -104.85
16.5531 | -104.8522
16.5568 | -104.8491
16.5604 | -104.8468
16.5641 | -104.8419
16.5677 | -104.8298
16.5714 | -104.8146
16.5751 | -104.7957
16.5787 | -104.775
16.5824 | -104.7528
16.586 | -104.7288
16.5897 | -104.6994
16.5934 | -104.6698
16.597 | -104.641
16.6007 | -104.6095
16.6043 | -104.5778
16.608 | -104.5473
16.6116 | -104.5176
16.6153 | -104.4858
16.619 | -104.4527
16.6226 | -104.4209
16.6263 | -104.3926
16.6299 | -104.3636
16.6336 | -104.3365
16.6372 | -104.3145
16.6409 | -104.294
16.6446 | -104.2785
16.6482 | -104.2641
16.6519 | -104.2524
16.6555 | -104.2451
16.6592 | -104.2368
16.6629 | -104.2301
16.6665 | -104.2311
16.6702 | -104.2358
16.6738 | -104.2425
16.6775 | -104.2543
16.6811 | -104.268
16.6848 | -104.2829
16.6885 | -104.2983
16.6921 | -104.317
16.6958 | -104.3392
16.6994 | -104.3613
16.7031 | -104.3832
16.7068 | -104.4055
16.7104 | -104.427
16.7141 | -104.448
16.7177 | -104.4671
16.7214 | -104.4851
16.725 | -104.5018
16.7287 | -104.5175
16.7324 | -104.533
16.736 | -104.5485
16.7397 | -104.5626
16.7433 | -104.5732
16.747 | -104.5785
16.7507 | -104.5798
16.7543 | -104.5818
16.758 | -104.5868
16.7616 | -104.5894
16.7653 | -104.5871
16.7689 | -104.5805
16.7726 | -104.5704
16.7763 | -104.561
16.7799 | -104.5516
16.7836 | -104.5368
16.7872 | -104.5203
16.7909 | -104.5037
16.7945 | -104.4877
16.7982 | -104.4693
16.8019 | -104.452
16.8055 | -104.4375
16.8092 | -104.4219
16.8128 | -104.4039
16.8165 | -104.3862
16.8202 | -104.3705
16.8238 | -104.3534
16.8275 | -104.3404
16.8311 | -104.3278
16.8348 | -104.3143
16.8384 | -104.3026
16.8421 | -104.293
16.8458 | -104.2806
16.8494 | -104.2669
16.8531 | -104.2585
16.8567 | -104.2499
16.8604 | -104.2413
16.8641 | -104.2314
16.8677 | -104.2246
16.8714 | -104.2185
16.875 | -104.2091
16.8787 | -104.201
16.8823 | -104.1897
16.886 | -104.1727
16.8897 | -104.1591
16.8933 | -104.1468
16.897 | -104.133
16.9006 | -104.118
16.9043 | -104.1025
16.908 | -104.0841
16.9116 | -104.0613
16.9153 | -104.0382
16.9189 | -104.0152
16.9226 | -103.99
16.9262 | -103.9655
16.9299 | -103.939
16.9336 | -103.9124
16.9372 | -103.8841
16.9409 | -103.8581
16.9445 | -103.8328
16.9482 | -103.8098
16.9519 | -103.7859
16.9555 | -103.762
16.9592 | -103.7406
16.9628 | -103.7186
16.9665 | -103.6981
16.9701 | -103.677
16.9738 | -103.655
16.9775 | -103.6371
16.9811 | -103.6198
16.9848 | -103.6042
16.9884 | -103.5928
16.9921 | -103.5848
16.9957 | -103.5763
16.9994 | -103.5682
17.0031 | -103.5635
17.0067 | -103.5603
17.0104 | -103.5603
17.014 | -103.5635
17.0177 | -103.5667
17.0214 | -103.5701
17.025 | -103.5777
17.0287 | -103.5862
17.0323 | -103.5932
17.036 | -103.6031
17.0396 | -103.6175
17.0433 | -103.632
17.047 | -103.6471
17.0506 | -103.665
17.0543 | -103.6856
17.0579 | -103.7037
17.0616 | -103.721
17.0653 | -103.7397
17.0689 | -103.7573
17.0726 | -103.7754
17.0762 | -103.7939
17.0799 | -103.8125
17.0835 | -103.8289
17.0872 | -103.8471
17.0909 | -103.8646
17.0945 | -103.8792
17.0982 | -103.8929
17.1018 | -103.9042
17.1055 | -103.9156
17.1091 | -103.9249
17.1128 | -103.9297
17.1165 | -103.9327
17.1201 | -103.936
17.1238 | -103.9401
17.1274 | -103.9457
17.1311 | -103.9508
17.1348 | -103.954
17.1384 | -103.956
17.1421 | -103.9586
17.1457 | -103.9594
17.1494 | -103.9592
17.153 | -103.9591
17.1567 | -103.9582
17.1604 | -103.9564
17.164 | -103.9545
17.1677 | -103.9533
17.1713 | -103.9532
17.175 | -103.954
17.1787 | -103.9523
17.1823 | -103.9501
17.186 | -103.9464
17.1896 | -103.9451
17.1933 | -103.9425
17.1969 | -103.9401
17.2006 | -103.9385
17.2043 | -103.9379
17.2079 | -103.9398
17.2116 | -103.9405
17.2152 | -103.9419
17.2189 | -103.9451
17.2226 | -103.9458
17.2262 | -103.9482
17.2299 | -103.9502
17.2335 | -103.9521
17.2372 | -103.9543
17.2408 | -103.956
17.2445 | -103.9608
17.2482 | -103.9648
17.2518 | -103.9689
17.2555 | -103.9741
17.2591 | -103.9766
17.2628 | -103.9797
17.2664 | -103.9783
17.2701 | -103.9763
17.2738 | -103.9783
17.2774 | -103.9785
17.2811 | -103.9765
17.2847 | -103.9721
17.2884 | -103.969
17.2921 | -103.9635
17.2957 | -103.9547
17.2994 | -103.9477
17.303 | -103.941
17.3067 | -103.9316
17.3103 | -103.9211
17.314 | -103.9081
17.3177 | -103.8947
17.3213 | -103.8801
17.325 | -103.861
17.3286 | -103.8428
17.3323 | -103.8243
17.336 | -103.8038
17.3396 | -103.7814
17.3433 | -103.7588
17.3469 | -103.7365
17.3506 | -103.7142
17.3542 | -103.6926
17.3579 | -103.6735
17.3616 | -103.6549
17.3652 | -103.6352
17.3689 | -103.6172
17.3725 | -103.5969
17.3762 | -103.5745
17.3799 | -103.5543
17.3835 | -103.5374
17.3872 | -103.5206
17.3908 | -103.5081
17.3945 | -103.4982
17.3981 | -103.4883
17.4018 | -103.4833
17.4055 | -103.4796
17.4091 | -103.4745
17.4128 | -103.4825
17.4164 | -103.4812
17.4201 | -103.4789
17.4238 | -103.4766
17.4274 | -103.48
17.4311 | -103.4845
17.4347 | -103.4896
17.4384 | -103.4938
17.442 | -103.5007
17.4457 | -103.5105
17.4494 | -103.5225
17.453 | -103.5342
17.4567 | -103.544
17.4603 | -103.5554
17.464 | -103.5676
17.4676 | -103.5814
17.4713 | -103.5972
17.475 | -103.613
17.4786 | -103.6287
17.4823 | -103.6432
17.4859 | -103.6584
17.4896 | -103.6725
17.4933 | -103.6846
17.4969 | -103.6987
17.5006 | -103.7135
17.5042 | -103.7284
17.5079 | -103.7429
17.5115 | -103.7559
17.5152 | -103.7742
17.5189 | -103.7897
17.5225 | -103.7988
17.5262 | -103.8102
17.5298 | -103.821
17.5335 | -103.8314
17.5372 | -103.8399
17.5408 | -103.8483
17.5445 | -103.8562
17.5481 | -103.8613
17.5518 | -103.8673
17.5554 | -103.872
17.5591 | -103.8725
17.5628 | -103.8741
17.5664 | -103.8738
17.5701 | -103.8725
17.5737 | -103.8732
17.5774 | -103.8712
17.5811 | -103.8677
17.5847 | -103.8666
17.5884 | -103.8633
17.592 | -103.8598
17.5957 | -103.8547
17.5993 | -103.8479
17.603 | -103.8404
17.6067 | -103.8318
17.6103 | -103.8244
17.614 | -103.8153
17.6176 | -103.8042
17.6213 | -103.7939
17.6249 | -103.7881
17.6286 | -103.782
17.6323 | -103.7718
17.6359 | -103.7601
17.6396 | -103.7471
17.6432 | -103.7334
17.6469 | -103.7198
17.6506 | -103.7061
17.6542 | -103.6907
17.6579 | -103.676
17.6615 | -103.6601
17.6652 | -103.6436
17.6688 | -103.6291
17.6725 | -103.6151
17.6762 | -103.6039
17.6798 | -103.5889
17.6835 | -103.5758
17.6871 | -103.5649
17.6908 | -103.5543
17.6945 | -103.5432
17.6981 | -103.5307
17.7018 | -103.5175
17.7054 | -103.5073
17.7091 | -103.4993
17.7127 | -103.4908
17.7164 | -103.4859
17.7201 | -103.4836
17.7237 | -103.4809
17.7274 | -103.4785
17.731 | -103.4779
17.7347 | -103.4764
17.7384 | -103.4742
17.742 | -103.4747
17.7457 | -103.4754
17.7493 | -103.4732
17.753 | -103.4722
17.7566 | -103.4698
17.7603 | -103.4675
17.764 | -103.465
17.7676 | -103.4598
17.7713 | -103.4526
17.7749 | -103.4443
17.7786 | -103.4366
17.7822 | -103.4276
17.7859 | -103.4184
17.7896 | -103.41
17.7932 | -103.3987
17.7969 | -103.3866
17.8005 | -103.3744
17.8042 | -103.3593
17.8079 | -103.3459
17.8115 | -103.3326
17.8152 | -103.317
17.8188 | -103.3013
17.8225 | -103.2843
17.8261 | -103.2664
17.8298 | -103.248
17.8335 | -103.2291
17.8371 | -103.2091
17.8408 | -103.1915
17.8444 | -103.1718
17.8481 | -103.1513
17.8518 | -103.1319
17.8554 | -103.1133
17.8591 | -103.0969
17.8627 | -103.0813
17.8664 | -103.0673
17.87 | -103.0561
17.8737 | -103.0468
17.8774 | -103.0357
17.881 | -103.0253
17.8847 | -103.017
17.8883 | -103.0124
17.892 | -103.0112
17.8957 | -103.0099
17.8993 | -103.0106
17.903 | -103.014
17.9066 | -103.0203
17.9103 | -103.0279
17.9139 | -103.0395
17.9176 | -103.0538
17.9213 | -103.0677
17.9249 | -103.0841
17.9286 | -103.1011
17.9322 | -103.1205
17.9359 | -103.1401
17.9395 | -103.161
17.9432 | -103.1849
17.9469 | -103.2057
17.9505 | -103.2261
17.9542 | -103.2495
17.9578 | -103.2736
17.9615 | -103.2947
17.9652 | -103.3161
17.9688 | -103.3391
17.9725 | -103.3581
17.9761 | -103.3738
17.9798 | -103.39
17.9834 | -103.4043
17.9871 | -103.4182
17.9908 | -103.4332
17.9944 | -103.4479
17.9981 | -103.4587
18.0017 | -103.466
18.0054 | -103.4754
18.0091 | -103.4841
18.0127 | -103.4925
18.0164 | -103.5002
18.02 | -103.5065
18.0237 | -103.5113
18.0273 | -103.514
18.031 | -103.517
18.0347 | -103.5157
18.0383 | -103.5116
18.042 | -103.5065
18.0456 | -103.5006
18.0493 | -103.4971
18.053 | -103.4915
18.0566 | -103.4841
18.0603 | -103.4766
18.0639 | -103.4676
18.0676 | -103.4576
18.0712 | -103.4474
18.0749 | -103.4363
18.0786 | -103.4251
18.0822 | -103.4123
18.0859 | -103.3992
18.0895 | -103.3865
18.0932 | -103.3698
18.0968 | -103.3517
18.1005 | -103.3368
18.1042 | -103.3219
18.1078 | -103.3064
18.1115 | -103.2915
18.1151 | -103.2777
18.1188 | -103.2623
18.1225 | -103.2469
18.1261 | -103.2312
18.1298 | -103.217
18.1334 | -103.2041
18.1371 | -103.193
18.1407 | -103.1826
18.1444 | -103.1717
18.1481 | -103.1587
18.1517 | -103.1498
18.1554 | -103.1419
18.159 | -103.1337
18.1627 | -103.1266
18.1664 | -103.12
18.17 | -103.11
18.1737 | -103.1012
18.1773 | -103.0934
18.181 | -103.087
18.1846 | -103.0811
18.1883 | -103.0743
18.192 | -103.069
18.1956 | -103.0631
18.1993 | -103.0579
18.2029 | -103.0522
18.2066 | -103.0458
18.2103 | -103.0424
18.2139 | -103.0379
18.2176 | -103.0322
18.2212 | -103.0282
18.2249 | -103.0243
18.2285 | -103.0196
18.2322 | -103.0154
18.2359 | -103.011
18.2395 | -103.0077
18.2432 | -103.0071
18.2468 | -103.0048
18.2505 | -103.0039
18.2541 | -103.0018
18.2578 | -102.999
18.2615 | -102.997
18.2651 | -102.9942
18.2688 | -102.9889
18.2724 | -102.9859
18.2761 | -102.9853
18.2798 | -102.9854
18.2834 | -102.9846
18.2871 | -102.9855
18.2907 | -102.9874
18.2944 | -102.9904
18.298 | -102.9938
18.3017 | -102.9952
18.3054 | -102.9965
18.309 | -102.9997
18.3127 | -103.0035
18.3163 | -103.006
18.32 | -103.0067
18.3237 | -103.0091
18.3273 | -103.0149
18.331 | -103.0186
18.3346 | -103.0224
18.3383 | -103.0253
18.3419 | -103.0278
18.3456 | -103.031
18.3493 | -103.0336
18.3529 | -103.0348
18.3566 | -103.0335
18.3602 | -103.0271
18.3639 | -103.0211
18.3676 | -103.0178
18.3712 | -103.011
18.3749 | -102.9984
18.3785 | -102.9863
18.3822 | -102.9726
18.3858 | -102.959
18.3895 | -102.9422
18.3932 | -102.9254
18.3968 | -102.9072
18.4005 | -102.8885
18.4041 | -102.869
18.4078 | -102.8488
18.4114 | -102.8251
18.4151 | -102.8004
18.4188 | -102.7736
18.4224 | -102.7468
18.4261 | -102.7213
18.4297 | -102.6953
18.4334 | -102.6696
18.4371 | -102.6432
18.4407 | -102.6158
18.4444 | -102.5909
18.448 | -102.569
18.4517 | -102.5454
18.4553 | -102.5226
18.459 | -102.5009
18.4627 | -102.4797
18.4663 | -102.4592
18.47 | -102.4494
18.4736 | -102.4313
18.4773 | -102.4154
18.481 | -102.4016
18.4846 | -102.3867
18.4883 | -102.373
18.4919 | -102.3614
18.4956 | -102.3505
18.4992 | -102.3435
18.5029 | -102.3361
18.5066 | -102.329
18.5102 | -102.3245
18.5139 | -102.3221
18.5175 | -102.322
18.5212 | -102.3235
18.5249 | -102.3266
18.5285 | -102.33
18.5322 | -102.3367
18.5358 | -102.342
18.5395 | -102.3485
18.5431 | -102.3569
18.5468 | -102.3675
18.5505 | -102.3794
18.5541 | -102.391
18.5578 | -102.4043
18.5614 | -102.4172
18.5651 | -102.4295
18.5688 | -102.4414
18.5724 | -102.4531
18.5761 | -102.4664
18.5797 | -102.4816
18.5834 | -102.4993
18.587 | -102.5133
18.5907 | -102.523
18.5944 | -102.5353
18.598 | -102.55
18.6017 | -102.5654
18.6053 | -102.581
18.609 | -102.5935
18.6126 | -102.6057
18.6163 | -102.6215
18.62 | -102.6372
18.6236 | -102.6545
18.6273 | -102.6705
18.6309 | -102.6807
18.6346 | -102.6915
18.6383 | -102.7039
18.6419 | -102.7149
18.6456 | -102.7258
18.6492 | -102.735
18.6529 | -102.7438
18.6565 | -102.7534
18.6602 | -102.7612
18.6639 | -102.7664
18.6675 | -102.7726
18.6712 | -102.7788
18.6748 | -102.782
18.6785 | -102.7846
18.6822 | -102.79
18.6858 | -102.7941
18.6895 | -102.7957
18.6931 | -102.7966
18.6968 | -102.7961
18.7004 | -102.795
18.7041 | -102.7953
18.7078 | -102.7923
18.7114 | -102.7896
18.7151 | -102.787
18.7187 | -102.7835
18.7224 | -102.7816
18.7261 | -102.7815
18.7297 | -102.7821
18.7334 | -102.7807
18.737 | -102.7798
18.7407 | -102.7793
18.7443 | -102.7769
18.748 | -102.7745
18.7517 | -102.7742
18.7553 | -102.7737
18.759 | -102.7735
18.7626 | -102.7726
18.7663 | -102.7672
18.7699 | -102.7582
18.7736 | -102.7521
18.7773 | -102.7479
18.7809 | -102.744
18.7846 | -102.7393
18.7882 | -102.7331
18.7919 | -102.7275
18.7956 | -102.7215
18.7992 | -102.7148
18.8029 | -102.7084
18.8065 | -102.6994
18.8102 | -102.687
18.8138 | -102.6726
18.8175 | -102.6586
18.8212 | -102.6453
18.8248 | -102.6317
18.8285 | -102.6204
18.8321 | -102.6086
18.8358 | -102.5972
18.8395 | -102.5846
18.8431 | -102.5692
18.8468 | -102.5487
18.8504 | -102.5294
18.8541 | -102.5136
18.8577 | -102.4991
18.8614 | -102.4827
18.8651 | -102.4644
18.8687 | -102.4466
18.8724 | -102.4299
18.876 | -102.4146
18.8797 | -102.3991
18.8834 | -102.3825
18.887 | -102.3665
18.8907 | -102.3502
18.8943 | -102.3338
18.898 | -102.3209
18.9016 | -102.3097
18.9053 | -102.3002
18.909 | -102.2926
18.9126 | -102.2858
18.9163 | -102.2789
18.9199 | -102.2741
18.9236 | -102.2712
18.9272 | -102.2707
18.9309 | -102.2722
18.9346 | -102.2733
18.9382 | -102.276
18.9419 | -102.2818
18.9455 | -102.2909
18.9492 | -102.2996
18.9529 | -102.3084
18.9565 | -102.3183
18.9602 | -102.3289
18.9638 | -102.3405
18.9675 | -102.3559
18.9711 | -102.3693
18.9748 | -102.3841
18.9785 | -102.3984
18.9821 | -102.4146
18.9858 | -102.4331
18.9894 | -102.4498
18.9931 | -102.466
18.9968 | -102.4821
19.0004 | -102.497
19.0041 | -102.5117
19.0077 | -102.5257
19.0114 | -102.5397
19.015 | -102.5525
19.0187 | -102.5654
19.0224 | -102.5789
19.026 | -102.5898
19.0297 | -102.5969
19.0333 | -102.6041
19.037 | -102.6136
19.0407 | -102.6233
19.0443 | -102.6267
19.048 | -102.6271
19.0516 | -102.627
19.0553 | -102.6262
19.0589 | -102.6258
19.0626 | -102.6239
19.0663 | -102.6197
19.0699 | -102.616
19.0736 | -102.6137
19.0772 | -102.611
19.0809 | -102.6085
19.0845 | -102.6019
19.0882 | -102.594
19.0919 | -102.5866
19.0955 | -102.5757
19.0992 | -102.5656
19.1028 | -102.5556
19.1065 | -102.5438
19.1102 | -102.53
19.1138 | -102.5166
19.1175 | -102.5047
19.1211 | -102.4913
19.1248 | -102.4751
19.1284 | -102.458
19.1321 | -102.4374
19.1358 | -102.4146
19.1394 | -102.392
19.1431 | -102.3708
19.1467 | -102.3498
19.1504 | -102.3263
19.1541 | -102.3034
19.1577 | -102.2803
19.1614 | -102.2581
19.165 | -102.2367
19.1687 | -102.2123
19.1723 | -102.1873
19.176 | -102.1614
19.1797 | -102.1348
19.1833 | -102.1091
19.187 | -102.0854
19.1906 | -102.062
19.1943 | -102.0357
19.198 | -102.0133
19.2016 | -101.9911
19.2053 | -101.9723
19.2089 | -101.955
19.2126 | -101.9385
19.2162 | -101.9226
19.2199 | -101.9084
19.2236 | -101.8971
19.2272 | -101.8898
19.2309 | -101.885
19.2345 | -101.8787
19.2382 | -101.8732
19.2418 | -101.8706
19.2455 | -101.8707
19.2492 | -101.8732
19.2528 | -101.8783
19.2565 | -101.8853
19.2601 | -101.8927
19.2638 | -101.9008
19.2675 | -101.9121
19.2711 | -101.9257
19.2748 | -101.9401
19.2784 | -101.9571
19.2821 | -101.9761
19.2857 | -101.9977
19.2894 | -102.0206
19.2931 | -102.043
19.2967 | -102.0637
19.3004 | -102.0854
19.304 | -102.109
19.3077 | -102.1346
19.3114 | -102.1587
19.315 | -102.1834
19.3187 | -102.2077
19.3223 | -102.2298
19.326 | -102.2472
19.3296 | -102.2668
19.3333 | -102.2849
19.337 | -102.3014
19.3406 | -102.3196
19.3443 | -102.3356
19.3479 | -102.3478
19.3516 | -102.3589
19.3553 | -102.369
19.3589 | -102.3808
19.3626 | -102.3895
19.3662 | -102.3935
19.3699 | -102.3962
19.3735 | -102.3957
19.3772 | -102.3972
19.3809 | -102.3981
19.3845 | -102.3986
19.3882 | -102.3986
19.3918 | -102.397
19.3955 | -102.392
19.3991 | -102.3861
19.4028 | -102.3791
19.4065 | -102.3746
19.4101 | -102.3702
19.4138 | -102.3648
19.4174 | -102.3564
19.4211 | -102.3485
19.4248 | -102.3422
19.4284 | -102.3346
19.4321 | -102.3277
19.4357 | -102.3195
19.4394 | -102.3097
19.443 | -102.301
19.4467 | -102.2943
19.4504 | -102.29
19.454 | -102.2872
19.4577 | -102.2856
19.4613 | -102.284
19.465 | -102.2832
19.4687 | -102.2856
19.4723 | -102.2848
19.476 | -102.2868
19.4796 | -102.2882
19.4833 | -102.2928
19.4869 | -102.2976
19.4906 | -102.3013
19.4943 | -102.3071
19.4979 | -102.3154
19.5016 | -102.3231
19.5052 | -102.3312
19.5089 | -102.3425
19.5126 | -102.3543
19.5162 | -102.3659
19.5199 | -102.3796
19.5235 | -102.397
19.5272 | -102.4088
19.5308 | -102.4214
19.5345 | -102.4351
19.5382 | -102.4499
19.5418 | -102.4662
19.5455 | -102.4798
19.5491 | -102.4931
19.5528 | -102.5072
19.5565 | -102.5212
19.5601 | -102.534
19.5638 | -102.5444
19.5674 | -102.5545
19.5711 | -102.5641
19.5747 | -102.5741
19.5784 | -102.5832
19.5821 | -102.5906
19.5857 | -102.5952
19.5894 | -102.5986
19.593 | -102.6023
19.5967 | -102.6037
19.6003 | -102.6031
19.604 | -102.603
19.6077 | -102.5978
19.6113 | -102.5934
19.615 | -102.5879
19.6186 | -102.5791
19.6223 | -102.5684
19.626 | -102.5603
19.6296 | -102.5475
19.6333 | -102.5318
19.6369 | -102.5182
19.6406 | -102.5044
19.6442 | -102.4872
19.6479 | -102.4704
19.6516 | -102.4504
19.6552 | -102.4258
19.6589 | -102.4013
19.6625 | -102.3775
19.6662 | -102.3512
19.6699 | -102.3265
19.6735 | -102.3029
19.6772 | -102.2789
19.6808 | -102.2526
19.6845 | -102.2244
19.6881 | -102.1953
19.6918 | -102.1681
19.6955 | -102.1423
19.6991 | -102.1185
19.7028 | -102.0939
19.7064 | -102.0695
19.7101 | -102.046
19.7138 | -102.0244
19.7174 | -102.0042
19.7211 | -101.9848
19.7247 | -101.9652
19.7284 | -101.9472
19.732 | -101.9308
19.7357 | -101.9157
19.7394 | -101.9031
19.743 | -101.8908
19.7467 | -101.8812
19.7503 | -101.8731
19.754 | -101.8667
19.7576 | -101.8618
19.7613 | -101.8568
19.765 | -101.8536
19.7686 | -101.8517
19.7723 | -101.8508
19.7759 | -101.8516
19.7796 | -101.8542
19.7833 | -101.8576
19.7869 | -101.8615
19.7906 | -101.8653
19.7942 | -101.8721
19.7979 | -101.8776
19.8015 | -101.8839
19.8052 | -101.8912
19.8089 | -101.898
19.8125 | -101.9057
19.8162 | -101.9115
19.8198 | -101.9197
19.8235 | -101.9287
19.8272 | -101.9381
19.8308 | -101.9453
19.8345 | -101.952
19.8381 | -101.959
19.8418 | -101.9658
19.8454 | -101.9739
19.8491 | -101.9834
19.8528 | -101.9913
19.8564 | -101.9973
19.8601 | -102.0048
19.8637 | -102.0118
19.8674 | -102.0184
19.8711 | -102.0249
19.8747 | -102.0318
19.8784 | -102.0404
19.882 | -102.05
19.8857 | -102.0613
19.8893 | -102.0738
19.893 | -102.0858
19.8967 | -102.0964
19.9003 | -102.1068
19.904 | -102.1171
19.9076 | -102.1261
19.9113 | -102.1383
19.9149 | -102.1508
19.9186 | -102.1627
19.9223 | -102.1737
19.9259 | -102.1873
19.9296 | -102.2003
19.9332 | -102.2143
19.9369 | -102.2297
19.9406 | -102.2454
19.9442 | -102.2574
19.9479 | -102.271
19.9515 | -102.2832
19.9552 | -102.2944
19.9588 | -102.3073
19.9625 | -102.3233
19.9662 | -102.3382
19.9698 | -102.3526
19.9735 | -102.3668
19.9771 | -102.3788
19.9808 | -102.3894
19.9845 | -102.4029
19.9881 | -102.4175
19.9918 | -102.4315
19.9954 | -102.4448
19.9991 | -102.4575
20.0027 | -102.4705
20.0064 | -102.4837
20.0101 | -102.4952
20.0137 | -102.506
20.0174 | -102.5182
20.021 | -102.53
20.0247 | -102.5428
20.0284 | -102.5561
20.032 | -102.5697
20.0357 | -102.5829
20.0393 | -102.5963
20.043 | -102.609
20.0466 | -102.6225
20.0503 | -102.6363
20.054 | -102.6498
20.0576 | -102.663
20.0613 | -102.6761
20.0649 | -102.6884
20.0686 | -102.7001
20.0722 | -102.712
20.0759 | -102.725
20.0796 | -102.7366
20.0832 | -102.7479
20.0869 | -102.76
20.0905 | -102.7729
20.0942 | -102.7883
20.0979 | -102.8026
20.1015 | -102.8195
20.1052 | -102.8338
20.1088 | -102.846
20.1125 | -102.8589
20.1161 | -102.876
20.1198 | -102.8927
20.1235 | -102.9073
20.1271 | -102.9237
20.1308 | -102.9378
20.1344 | -102.9521
20.1381 | -102.9679
20.1418 | -102.9846
20.1454 | -102.9991
20.1491 | -103.0124
20.1527 | -103.025
20.1564 | -103.0374
20.16 | -103.0489
20.1637 | -103.0591
20.1674 | -103.0675
20.171 | -103.0729
20.1747 | -103.078
20.1783 | -103.0817
20.182 | -103.0881
20.1857 | -103.0944
20.1893 | -103.099
20.193 | -103.1031
20.1966 | -103.1072
20.2003 | -103.1103
20.2039 | -103.1099
20.2076 | -103.1079
20.2113 | -103.1065
20.2149 | -103.104
20.2186 | -103.1021
20.2222 | -103.0977
20.2259 | -103.0909
20.2295 | -103.0859
20.2332 | -103.0796
20.2369 | -103.0722
20.2405 | -103.066
20.2442 | -103.0585
20.2478 | -103.0503
20.2515 | -103.0449
20.2552 | -103.0392
20.2588 | -103.0302
20.2625 | -103.0226
20.2661 | -103.0154
20.2698 | -103.0076
20.2734 | -103.0004
20.2771 | -102.9947
20.2808 | -102.9903
20.2844 | -102.986
20.2881 | -102.9818
20.2917 | -102.9786
20.2954 | -102.9759
20.2991 | -102.9726
20.3027 | -102.967
20.3064 | -102.9632
20.31 | -102.9618
20.3137 | -102.9588
20.3173 | -102.9578
20.321 | -102.9571
20.3247 | -102.9568
20.3283 | -102.956
20.332 | -102.9564
20.3356 | -102.9559
20.3393 | -102.9547
20.343 | -102.9546
20.3466 | -102.9564
20.3503 | -102.9566
20.3539 | -102.9591
20.3576 | -102.9624
20.3612 | -102.9653
20.3649 | -102.9663
20.3686 | -102.9668
20.3722 | -102.9689
20.3759 | -102.9688
20.3795 | -102.9686
20.3832 | -102.9681
20.3868 | -102.9679
20.3905 | -102.9654
20.3942 | -102.9648
20.3978 | -102.965
20.4015 | -102.9644
20.4051 | -102.9637
20.4088 | -102.9627
20.4125 | -102.9602
20.4161 | -102.9576
20.4198 | -102.9572
20.4234 | -102.9549
20.4271 | -102.954
20.4307 | -102.9523
20.4344 | -102.9472
20.4381 | -102.9432
20.4417 | -102.9383
20.4454 | -102.9352
20.449 | -102.9308
20.4527 | -102.927
20.4564 | -102.9213
20.46 | -102.9117
20.4637 | -102.902
20.4673 | -102.8932
20.471 | -102.8842
20.4746 | -102.8752
20.4783 | -102.8664
20.482 | -102.8567
20.4856 | -102.8463
20.4893 | -102.836
20.4929 | -102.8222
20.4966 | -102.8087
20.5003 | -102.796
20.5039 | -102.7838
20.5076 | -102.7692
20.5112 | -102.7555
20.5149 | -102.741
20.5185 | -102.7256
20.5222 | -102.7097
20.5259 | -102.692
20.5295 | -102.6705
20.5332 | -102.6478
20.5368 | -102.6282
20.5405 | -102.6079
20.5441 | -102.5869
20.5478 | -102.5628
20.5515 | -102.5411
20.5551 | -102.5196
20.5588 | -102.494
20.5624 | -102.466
20.5661 | -102.4375
20.5698 | -102.411
20.5734 | -102.3858
20.5771 | -102.3595
20.5807 | -102.3281
20.5844 | -102.3
20.588 | -102.2729
20.5917 | -102.245
20.5954 | -102.2131
20.599 | -102.1801
20.6027 | -102.1507
20.6063 | -102.1198
20.61 | -102.0843
20.6137 | -102.0485
20.6173 | -102.0182
20.621 | -101.9893
20.6246 | -101.9579
20.6283 | -101.926
20.6319 | -101.8915
20.6356 | -101.857
20.6393 | -101.8232
20.6429 | -101.7903
20.6466 | -101.7561
20.6502 | -101.7244
20.6539 | -101.6913
20.6576 | -101.6596
20.6612 | -101.6297
20.6649 | -101.5969
20.6685 | -101.5651
20.6722 | -101.5352
20.6758 | -101.5053
20.6795 | -101.4753
20.6832 | -101.4461
20.6868 | -101.4162
20.6905 | -101.3873
20.6941 | -101.3568
20.6978 | -101.3262
20.7015 | -101.2977
20.7051 | -101.2712
20.7088 | -101.2433
20.7124 | -101.2171
20.7161 | -101.1931
20.7197 | -101.1686
20.7234 | -101.1442
20.7271 | -101.1217
20.7307 | -101.0961
20.7344 | -101.0718
20.738 | -101.0499
20.7417 | -101.029
20.7453 | -101.0078
20.749 | -100.9884
20.7527 | -100.9704
20.7563 | -100.9502
20.76 | -100.9303
20.7636 | -100.9103
20.7673 | -100.8896
20.771 | -100.8711
20.7746 | -100.8528
20.7783 | -100.8347
20.7819 | -100.8181
20.7856 | -100.8046
20.7892 | -100.7918
20.7929 | -100.7779
20.7966 | -100.7636
20.8002 | -100.7494
20.8039 | -100.7353
20.8075 | -100.722
20.8112 | -100.7106
20.8149 | -100.6982
20.8185 | -100.6861
20.8222 | -100.6759
20.8258 | -100.667
20.8295 | -100.6551
20.8331 | -100.6442
20.8368 | -100.6349
20.8405 | -100.6258
20.8441 | -100.6133
20.8478 | -100.604
20.8514 | -100.5969
20.8551 | -100.5912
20.8588 | -100.5861
20.8624 | -100.579
20.8661 | -100.5727
20.8697 | -100.565
20.8734 | -100.5583
20.877 | -100.5541
20.8807 | -100.548
20.8844 | -100.5405
20.888 | -100.5361
20.8917 | -100.5356
20.8953 | -100.5336
20.899 | -100.5286
20.9026 | -100.5245
20.9063 | -100.5221
20.91 | -100.52
20.9136 | -100.5118
20.9173 | -100.5036
20.9209 | -100.4984
20.9246 | -100.494
20.9283 | -100.4898
20.9319 | -100.4862
20.9356 | -100.4831
20.9392 | -100.479
20.9429 | -100.4749
20.9465 | -100.4707
20.9502 | -100.467
20.9539 | -100.4628
20.9575 | -100.4582
20.9612 | -100.456
20.9648 | -100.4522
20.9685 | -100.4473
20.9722 | -100.4432
20.9758 | -100.4405
20.9795 | -100.4381
20.9831 | -100.4369
20.9868 | -100.4363
20.9904 | -100.4363
20.9941 | -100.4359
20.9978 | -100.4357
21.0014 | -100.4361
21.0051 | -100.4383
21.0087 | -100.4395
21.0124 | -100.44
21.016 | -100.442
21.0197 | -100.4453
21.0234 | -100.4498
21.027 | -100.4578
21.0307 | -100.4651
21.0343 | -100.4707
21.038 | -100.4762
21.0417 | -100.4823
21.0453 | -100.4905
21.049 | -100.5007
21.0526 | -100.5128
21.0563 | -100.5241
21.0599 | -100.5351
21.0636 | -100.5466
21.0673 | -100.5562
21.0709 | -100.5669
21.0746 | -100.58
21.0782 | -100.5915
21.0819 | -100.6018
21.0856 | -100.612
21.0892 | -100.6233
21.0929 | -100.6373
21.0965 | -100.6514
21.1002 | -100.6628
21.1038 | -100.6733
21.1075 | -100.6803
21.1112 | -100.69
21.1148 | -100.7013
21.1185 | -100.7136
21.1221 | -100.7244
21.1258 | -100.7343
21.1295 | -100.7454
21.1331 | -100.7554
21.1368 | -100.7649
21.1404 | -100.7752
21.1441 | -100.7858
21.1477 | -100.7932
21.1514 | -100.7992
21.1551 | -100.8039
21.1587 | -100.8058
21.1624 | -100.8093
21.166 | -100.8122
21.1697 | -100.817
21.1733 | -100.8199
21.177 | -100.8219
21.1807 | -100.8286
21.1843 | -100.8355
21.188 | -100.8386
21.1916 | -100.8408
21.1953 | -100.8446
21.199 | -100.8485
21.2026 | -100.852
21.2063 | -100.857
21.2099 | -100.8604
21.2136 | -100.8639
21.2172 | -100.8657
21.2209 | -100.8663
21.2246 | -100.8651
21.2282 | -100.8643
21.2319 | -100.861
21.2355 | -100.8532
21.2392 | -100.848
21.2429 | -100.8433
21.2465 | -100.8384
21.2502 | -100.8328
21.2538 | -100.8257
21.2575 | -100.8187
21.2611 | -100.812
21.2648 | -100.8062
21.2685 | -100.8007
21.2721 | -100.7928
21.2758 | -100.7818
21.2794 | -100.7716
21.2831 | -100.7619
21.2868 | -100.752
21.2904 | -100.7438
21.2941 | -100.7347
21.2977 | -100.7229
21.3014 | -100.7105
21.305 | -100.6993
21.3087 | -100.689
21.3124 | -100.6791
21.316 | -100.6698
21.3197 | -100.6607
21.3233 | -100.6516
21.327 | -100.6422
21.3307 | -100.6339
21.3343 | -100.6261
21.338 | -100.6198
21.3416 | -100.6165
21.3453 | -100.6125
21.3489 | -100.6065
21.3526 | -100.5983
21.3563 | -100.5923
21.3599 | -100.5878
21.3636 | -100.5839
21.3672 | -100.5808
21.3709 | -100.5769
21.3745 | -100.5711
21.3782 | -100.5665
21.3819 | -100.5615
21.3855 | -100.555
21.3892 | -100.5484
21.3928 | -100.5424
21.3965 | -100.5363
21.4002 | -100.5318
21.4038 | -100.5258
21.4075 | -100.5163
21.4111 | -100.506
21.4148 | -100.4963
21.4184 | -100.4862
21.4221 | -100.4727
21.4258 | -100.4572
21.4294 | -100.4389
21.4331 | -100.418
21.4367 | -100.3956
21.4404 | -100.3675
21.4441 | -100.3414
21.4477 | -100.3174
21.4514 | -100.2929
21.455 | -100.2676
21.4587 | -100.2376
21.4623 | -100.21
21.466 | -100.1812
21.4697 | -100.1471
21.4733 | -100.1083
21.477 | -100.0715
21.4806 | -100.0337
21.4843 | -99.9913
21.488 | -99.952
21.4916 | -99.9122
21.4953 | -99.8718
21.4989 | -99.8301
21.5026 | -99.7867
21.5062 | -99.7439
21.5099 | -99.6997
21.5136 | -99.6569
21.5172 | -99.616
21.5209 | -99.5763
21.5245 | -99.5359
21.5282 | -99.4963
21.5318 | -99.4569
21.5355 | -99.4196
21.5392 | -99.384
21.5428 | -99.3471
21.5465 | -99.311
21.5501 | -99.2771
21.5538 | -99.2464
21.5575 | -99.2178
21.5611 | -99.1891
21.5648 | -99.1599
21.5684 | -99.1335
21.5721 | -99.1097
21.5757 | -99.0879
21.5794 | -99.0683
21.5831 | -99.0513
21.5867 | -99.0368
21.5904 | -99.0242
21.594 | -99.0129
21.5977 | -99.0037
21.6014 | -98.9971
21.605 | -98.9941
21.6087 | -98.993
21.6123 | -98.9949
21.616 | -98.9977
21.6196 | -99.0004
21.6233 | -99.0063
21.627 | -99.0131
21.6306 | -99.0223
21.6343 | -99.0324
21.6379 | -99.0528
21.6416 | -99.063
21.6453 | -99.0738
21.6489 | -99.0868
21.6526 | -99.0988
21.6562 | -99.113
21.6599 | -99.132
21.6635 | -99.15
21.6672 | -99.1675
21.6709 | -99.1852
21.6745 | -99.2043
21.6782 | -99.2234
21.6818 | -99.2424
21.6855 | -99.2608
21.6891 | -99.2803
21.6928 | -99.3007
21.6965 | -99.3181
21.7001 | -99.3345
21.7038 | -99.351
21.7074 | -99.369
21.7111 | -99.3892
21.7148 | -99.4102
21.7184 | -99.4291
21.7221 | -99.4467
21.7257 | -99.4627
21.7294 | -99.479
21.733 | -99.4963
21.7367 | -99.5143
21.7404 | -99.5335
21.744 | -99.554
21.7477 | -99.5765
21.7513 | -99.5952
21.755 | -99.6092
21.7587 | -99.6221
21.7623 | -99.6365
21.766 | -99.6537
21.7696 | -99.6677
21.7733 | -99.6793
21.7769 | -99.6923
21.7806 | -99.7078
21.7843 | -99.722
21.7879 | -99.7339
21.7916 | -99.7469
21.7952 | -99.7562
21.7989 | -99.7644
21.8026 | -99.7691
21.8062 | -99.774
21.8099 | -99.7819
21.8135 | -99.7876
21.8172 | -99.7937
21.8208 | -99.7992
21.8245 | -99.8033
21.8282 | -99.8063
21.8318 | -99.8102
21.8355 | -99.813
21.8391 | -99.8172
21.8428 | -99.8179
21.8464 | -99.8197
21.8501 | -99.8224
21.8538 | -99.8243
21.8574 | -99.8241
21.8611 | -99.8229
21.8647 | -99.821
21.8684 | -99.819
21.8721 | -99.8182
21.8757 | -99.8174
21.8794 | -99.8152
21.883 | -99.8132
21.8867 | -99.8123
21.8903 | -99.812
21.894 | -99.8101
21.8977 | -99.807
21.9013 | -99.8053
21.905 | -99.8036
21.9086 | -99.8025
21.9123 | -99.7997
21.916 | -99.7966
21.9196 | -99.7934
21.9233 | -99.7929
21.9269 | -99.794
21.9306 | -99.7937
21.9342 | -99.7952
21.9379 | -99.7968
21.9416 | -99.7992
21.9452 | -99.8024
21.9489 | -99.8067
21.9525 | -99.8114
21.9562 | -99.817
21.9599 | -99.8191
21.9635 | -99.8257
21.9672 | -99.8313
21.9708 | -99.834
21.9745 | -99.8391
21.9781 | -99.8466
21.9818 | -99.8549
21.9855 | -99.8625
21.9891 | -99.8708
21.9928 | -99.88
21.9964 | -99.8896
22.0001 | -99.8958
22.0038 | -99.8999
22.0074 | -99.9092
22.0111 | -99.9196
22.0147 | -99.9299
22.0184 | -99.9387
22.022 | -99.9454
22.0257 | -99.9505
22.0294 | -99.9557
22.033 | -99.9611
22.0367 | -99.9688
22.0403 | -99.9761
22.044 | -99.982
22.0476 | -99.9863
22.0513 | -99.9905
22.055 | -99.9928
22.0586 | -99.9953
22.0623 | -99.9968
22.0659 | -99.9988
22.0696 | -99.9998
22.0733 | -99.9992
22.0769 | -99.9986
22.0806 | -99.997
22.0842 | -99.9959
22.0879 | -99.9956
22.0915 | -99.9937
22.0952 | -99.991
22.0989 | -99.9894
22.1025 | -99.9868
22.1062 | -99.984
22.1098 | -99.9849
22.1135 | -99.9899
22.1172 | -99.9946
22.1208 | -99.9956
22.1245 | -99.9978
22.1281 | -99.9984
22.1318 | -99.9991
22.1354 | -100.0006
22.1391 | -100.0013
22.1428 | -100.0044
22.1464 | -100.0104
22.1501 | -100.0148
22.1537 | -100.02
22.1574 | -100.0259
22.161 | -100.0331
22.1647 | -100.0477
22.1684 | -100.0571
22.172 | -100.0674
22.1757 | -100.0776
22.1793 | -100.0884
22.183 | -100.1023
22.1867 | -100.1165
22.1903 | -100.1294
22.194 | -100.1446
22.1976 | -100.1609
22.2013 | -100.1761
22.2049 | -100.1907
22.2086 | -100.2073
22.2123 | -100.2257
22.2159 | -100.2439
22.2196 | -100.2646
22.2232 | -100.2854
22.2269 | -100.3055
22.2306 | -100.3259
22.2342 | -100.3477
22.2379 | -100.3697
22.2415 | -100.3934
22.2452 | -100.418
22.2488 | -100.4422
22.2525 | -100.4673
22.2562 | -100.4926
22.2598 | -100.5177
22.2635 | -100.5439
22.2671 | -100.5721
22.2708 | -100.6004
22.2745 | -100.6309
22.2781 | -100.6596
22.2818 | -100.6907
22.2854 | -100.7222
22.2891 | -100.7517
22.2927 | -100.7801
22.2964 | -100.8076
22.3001 | -100.8369
22.3037 | -100.8674
22.3074 | -100.8975
22.311 | -100.9289
22.3147 | -100.9638
22.3183 | -101.0003
22.322 | -101.0328
22.3257 | -101.0636
22.3293 | -101.0947
22.333 | -101.1268
22.3366 | -101.1581
22.3403 | -101.1911
22.344 | -101.2253
22.3476 | -101.259
22.3513 | -101.2898
22.3549 | -101.3212
22.3586 | -101.3538
22.3622 | -101.3868
22.3659 | -101.4167
22.3696 | -101.4459
22.3732 | -101.4762
22.3769 | -101.5093
22.3805 | -101.543
22.3842 | -101.5762
22.3879 | -101.6078
22.3915 | -101.6378
22.3952 | -101.6693
22.3988 | -101.7002
22.4025 | -101.7308
22.4061 | -101.7612
22.4098 | -101.789
22.4135 | -101.8167
22.4171 | -101.8442
22.4208 | -101.8725
22.4244 | -101.897
22.4281 | -101.9196
22.4318 | -101.9385
22.4354 | -101.9556
22.4391 | -101.9748
22.4427 | -101.996
22.4464 | -102.0142
22.45 | -102.0321
22.4537 | -102.0517
22.4574 | -102.0609
22.461 | -102.0628
22.4647 | -102.0667
22.4683 | -102.0675
22.472 | -102.0645
22.4757 | -102.0645
22.4793 | -102.0657
22.483 | -102.0679
22.4866 | -102.0679
22.4903 | -102.067
22.4939 | -102.0658
22.4976 | -102.0625
22.5013 | -102.0554
22.5049 | -102.048
22.5086 | -102.0394
22.5122 | -102.0262
22.5159 | -102.0117
22.5195 | -101.9973
22.5232 | -101.9822
22.5269 | -101.9648
22.5305 | -101.9477
22.5342 | -101.9305
22.5378 | -101.9121
22.5415 | -101.8942
22.5452 | -101.8767
22.5488 | -101.8579
22.5525 | -101.8382
22.5561 | -101.8188
22.5598 | -101.7992
22.5634 | -101.7803
22.5671 | -101.7616
22.5708 | -101.7418
22.5744 | -101.7261
22.5781 | -101.7135
22.5817 | -101.7009
22.5854 | -101.6901
22.5891 | -101.6819
22.5927 | -101.6749
22.5964 | -101.6704
22.6 | -101.6676
22.6037 | -101.6676
22.6073 | -101.6701
22.611 | -101.676
22.6147 | -101.6859
22.6183 | -101.695
22.622 | -101.7054
22.6256 | -101.7205
22.6293 | -101.7385
22.633 | -101.7579
22.6366 | -101.7799
22.6403 | -101.8048
22.6439 | -101.8307
22.6476 | -101.8581
22.6512 | -101.8881
22.6549 | -101.9237
22.6586 | -101.9614
22.6622 | -101.999
22.6659 | -102.0365
22.6695 | -102.0762
22.6732 | -102.1163
22.6768 | -102.1563
22.6805 | -102.1956
22.6842 | -102.2372
22.6878 | -102.2792
22.6915 | -102.3127
22.6951 | -102.3503
22.6988 | -102.3861
22.7025 | -102.4225
22.7061 | -102.4574
22.7098 | -102.4912
22.7134 | -102.5247
22.7171 | -102.5544
22.7207 | -102.5832
22.7244 | -102.6099
22.7281 | -102.6332
22.7317 | -102.6541
22.7354 | -102.6722
22.739 | -102.6903
22.7427 | -102.7065
22.7464 | -102.7197
22.75 | -102.7299
22.7537 | -102.7405
22.7573 | -102.7509
22.761 | -102.7567
22.7646 | -102.7599
22.7683 | -102.7623
22.772 | -102.7621
22.7756 | -102.7585
22.7793 | -102.7564
22.7829 | -102.7561
22.7866 | -102.7502
22.7903 | -102.7461
22.7939 | -102.7389
22.7976 | -102.7228
22.8012 | -102.7023
22.8049 | -102.6806
22.8085 | -102.6587
22.8122 | -102.638
22.8159 | -102.6177
22.8195 | -102.5993
22.8232 | -102.5793
22.8268 | -102.5589
22.8305 | -102.5386
22.8341 | -102.5162
22.8378 | -102.4972
22.8415 | -102.4769
22.8451 | -102.4535
22.8488 | -102.4329
22.8524 | -102.4133
22.8561 | -102.3931
22.8598 | -102.3716
22.8634 | -102.3523
22.8671 | -102.334
22.8707 | -102.3164
22.8744 | -102.2995
22.878 | -102.2807
22.8817 | -102.2625
22.8854 | -102.2463
22.889 | -102.2308
22.8927 | -102.2162
22.8963 | -102.1993
22.9 | -102.1801
22.9037 | -102.1642
22.9073 | -102.1469
22.911 | -102.1274
22.9146 | -102.1104
22.9183 | -102.0954
22.9219 | -102.0814
22.9256 | -102.0686
22.9293 | -102.0575
22.9329 | -102.0441
22.9366 | -102.0321
22.9402 | -102.0227
22.9439 | -102.0122
22.9476 | -102.0018
22.9512 | -101.9923
22.9549 | -101.9843
22.9585 | -101.9762
22.9622 | -101.9668
22.9658 | -101.9575
22.9695 | -101.9502
22.9732 | -101.9429
22.9768 | -101.9338
22.9805 | -101.9254
22.9841 | -101.9198
22.9878 | -101.9155
22.9914 | -101.9115
22.9951 | -101.9084
22.9988 | -101.9026
23.0024 | -101.8975
23.0061 | -101.8928
23.0097 | -101.8896
23.0134 | -101.8867
23.0171 | -101.8843
23.0207 | -101.882
23.0244 | -101.8798
23.028 | -101.8763
23.0317 | -101.873
23.0353 | -101.8713
23.039 | -101.8689
23.0427 | -101.8674
23.0463 | -101.8657
23.05 | -101.8632
23.0536 | -101.862
23.0573 | -101.8603
23.061 | -101.8573
23.0646 | -101.8545
23.0683 | -101.8512
23.0719 | -101.8471
23.0756 | -101.8445
23.0792 | -101.8426
23.0829 | -101.8401
23.0866 | -101.8364
23.0902 | -101.8317
23.0939 | -101.8271
23.0975 | -101.824
23.1012 | -101.822
23.1049 | -101.8198
23.1085 | -101.8184
23.1122 | -101.8173
23.1158 | -101.816
23.1195 | -101.8157
23.1231 | -101.8149
23.1268 | -101.8148
23.1305 | -101.8165
23.1341 | -101.8177
23.1378 | -101.8187
23.1414 | -101.8206
23.1451 | -101.8249
23.1487 | -101.8312
23.1524 | -101.8379
23.1561 | -101.8463
23.1597 | -101.8584
23.1634 | -101.8689
23.167 | -101.881
23.1707 | -101.8956
23.1744 | -101.9095
23.178 | -101.9237
23.1817 | -101.9388
23.1853 | -101.9553
23.189 | -101.9709
23.1926 | -101.9886
23.1963 | -102.0081
23.2 | -102.0271
23.2036 | -102.0482
23.2073 | -102.07
23.2109 | -102.0924
23.2146 | -102.1139
23.2183 | -102.1392
23.2219 | -102.1624
23.2256 | -102.1842
23.2292 | -102.208
23.2329 | -102.2294
23.2365 | -102.254
23.2402 | -102.2799
23.2439 | -102.3032
23.2475 | -102.3238
23.2512 | -102.343
23.2548 | -102.3627
23.2585 | -102.3819
23.2622 | -102.402
23.2658 | -102.4206
23.2695 | -102.4361
23.2731 | -102.4545
23.2768 | -102.4729
23.2804 | -102.4872
23.2841 | -102.5019
23.2878 | -102.5183
23.2914 | -102.5352
23.2951 | -102.5511
23.2987 | -102.564
23.3024 | -102.5756
23.306 | -102.5892
23.3097 | -102.6035
23.3134 | -102.6138
23.317 | -102.623
23.3207 | -102.6301
23.3243 | -102.6365
23.328 | -102.647
23.3317 | -102.6563
23.3353 | -102.6633
23.339 | -102.6684
23.3426 | -102.6721
23.3463 | -102.6752
23.3499 | -102.6795
23.3536 | -102.6829
23.3573 | -102.6858
23.3609 | -102.689
23.3646 | -102.6894
23.3682 | -102.6899
23.3719 | -102.6913
23.3756 | -102.6924
23.3792 | -102.6946
23.3829 | -102.6951
23.3865 | -102.6949
23.3902 | -102.6977
23.3938 | -102.6982
23.3975 | -102.699
23.4012 | -102.7026
23.4048 | -102.7056
23.4085 | -102.7072
23.4121 | -102.7094
23.4158 | -102.7109
23.4195 | -102.7125
23.4231 | -102.7131
23.4268 | -102.7124
23.4304 | -102.7126
23.4341 | -102.7134
23.4377 | -102.7144
23.4414 | -102.7155
23.4451 | -102.7151
23.4487 | -102.7157
23.4524 | -102.7171
23.456 | -102.7181
23.4597 | -102.7175
23.4633 | -102.7183
23.467 | -102.7196
23.4707 | -102.7213
23.4743 | -102.7199
23.478 | -102.7204
23.4816 | -102.722
23.4853 | -102.7223
23.489 | -102.7226
23.4926 | -102.723
23.4963 | -102.7218
23.4999 | -102.7221
23.5036 | -102.7217
23.5072 | -102.7205
23.5109 | -102.7198
23.5146 | -102.7196
23.5182 | -102.7172
23.5219 | -102.7132
23.5255 | -102.709
23.5292 | -102.706
23.5329 | -102.703
23.5365 | -102.6993
23.5402 | -102.6924
23.5438 | -102.6883
23.5475 | -102.6823
23.5511 | -102.6767
23.5548 | -102.6701
23.5585 | -102.666
23.5621 | -102.6607
23.5658 | -102.6555
23.5694 | -102.6492
23.5731 | -102.6421
23.5768 | -102.6345
23.5804 | -102.6275
23.5841 | -102.6216
23.5877 | -102.6164
23.5914 | -102.6111
23.595 | -102.6067
23.5987 | -102.6052
23.6024 | -102.6068
23.606 | -102.6047
23.6097 | -102.5989
23.6133 | -102.5966
23.617 | -102.5943
23.6207 | -102.5918
23.6243 | -102.5927
23.628 | -102.5938
23.6316 | -102.5977
23.6353 | -102.601
23.6389 | -102.6035
23.6426 | -102.607
23.6463 | -102.6123
23.6499 | -102.6182
23.6536 | -102.6263
23.6572 | -102.634
23.6609 | -102.6411
23.6645 | -102.6489
23.6682 | -102.658
23.6719 | -102.6678
23.6755 | -102.6779
23.6792 | -102.6897
23.6828 | -102.6999
23.6865 | -102.7101
23.6902 | -102.7202
23.6938 | -102.7301
23.6975 | -102.7413
23.7011 | -102.7519
23.7048 | -102.7634
23.7084 | -102.7737
23.7121 | -102.7831
23.7158 | -102.7941
23.7194 | -102.8049
23.7231 | -102.813
23.7267 | -102.8223
23.7304 | -102.8296
23.7341 | -102.8357
23.7377 | -102.8403
23.7414 | -102.848
23.745 | -102.8555
23.7487 | -102.8619
23.7523 | -102.8718
23.756 | -102.8819
23.7597 | -102.8924
23.7633 | -102.9011
23.767 | -102.908
23.7706 | -102.9138
23.7743 | -102.9194
23.778 | -102.9266
23.7816 | -102.9341
23.7853 | -102.939
23.7889 | -102.9439
23.7926 | -102.9501
23.7962 | -102.9552
23.7999 | -102.959
23.8036 | -102.9602
23.8072 | -102.9619
23.8109 | -102.9668
23.8145 | -102.9692
23.8182 | -102.9714
23.8218 | -102.9745
23.8255 | -102.9773
23.8292 | -102.9802
23.8328 | -102.9841
23.8365 | -102.9863
23.8401 | -102.9895
23.8438 | -102.9931
23.8475 | -102.9969
23.8511 | -103.001
23.8548 | -103.0069
23.8584 | -103.0111
23.8621 | -103.014
23.8657 | -103.0173
23.8694 | -103.0222
23.8731 | -103.0261
23.8767 | -103.0299
23.8804 | -103.0314
23.884 | -103.0337
23.8877 | -103.0348
23.8914 | -103.0387
23.895 | -103.0417
23.8987 | -103.0439
23.9023 | -103.0463
23.906 | -103.0477
23.9096 | -103.0494
23.9133 | -103.0529
23.917 | -103.056
23.9206 | -103.0561
23.9243 | -103.057
23.9279 | -103.0577
23.9316 | -103.0585
23.9353 | -103.0577
23.9389 | -103.0578
23.9426 | -103.0597
23.9462 | -103.064
23.9499 | -103.0657
23.9535 | -103.0667
23.9572 | -103.0687
23.9609 | -103.0676
23.9645 | -103.0653
23.9682 | -103.0614
23.9718 | -103.0566
23.9755 | -103.0541
23.9791 | -103.0491
23.9828 | -103.0437
23.9865 | -103.0403
23.9901 | -103.0345
23.9938 | -103.0266
23.9974 | -103.018
24.0011 | -103.0097
24.0048 | -102.9989
24.0084 | -102.9861
24.0121 | -102.9734
24.0157 | -102.9598
24.0194 | -102.9463
24.023 | -102.9321
24.0267 | -102.9183
24.0304 | -102.9047
24.034 | -102.8903
24.0377 | -102.8735
24.0413 | -102.8528
24.045 | -102.8327
24.0487 | -102.8148
24.0523 | -102.7971
24.056 | -102.7788
24.0596 | -102.7593
24.0633 | -102.7384
24.0669 | -102.7211
24.0706 | -102.7038
24.0743 | -102.6847
24.0779 | -102.6622
24.0816 | -102.6407
24.0852 | -102.6193
24.0889 | -102.5988
24.0926 | -102.5771
24.0962 | -102.5563
24.0999 | -102.5362
24.1035 | -102.5178
24.1072 | -102.4992
24.1108 | -102.484
24.1145 | -102.4676
24.1182 | -102.4511
24.1218 | -102.4354
24.1255 | -102.4198
24.1291 | -102.4045
24.1328 | -102.3902
24.1364 | -102.3776
24.1401 | -102.3665
24.1438 | -102.3578
24.1474 | -102.3485
24.1511 | -102.3391
24.1547 | -102.332
24.1584 | -102.3253
24.1621 | -102.3196
24.1657 | -102.3143
24.1694 | -102.3091
24.173 | -102.3057
24.1767 | -102.3021
24.1803 | -102.2982
24.184 | -102.2961
24.1877 | -102.2944
24.1913 | -102.2927
24.195 | -102.2919
24.1986 | -102.2946
24.2023 | -102.2986
24.206 | -102.2999
24.2096 | -102.3019
24.2133 | -102.3049
24.2169 | -102.3106
24.2206 | -102.3172
24.2242 | -102.3245
24.2279 | -102.3318
24.2316 | -102.3385
24.2352 | -102.3455
24.2389 | -102.354
24.2425 | -102.3617
24.2462 | -102.3699
24.2499 | -102.3777
24.2535 | -102.3859
24.2572 | -102.3932
24.2608 | -102.3997
24.2645 | -102.4015
24.2681 | -102.4008
24.2718 | -102.3922
24.2755 | -102.3377
24.2791 | -102.3128
24.2828 | -102.2852
24.2864 | -102.2575
24.2901 | -102.2289
24.2937 | -102.2042
24.2974 | -102.1816
24.3011 | -102.1552
24.3047 | -102.1275
24.3084 | -102.098
24.312 | -102.0663
24.3157 | -102.0331
24.3194 | -101.9945
24.323 | -101.9536
24.3267 | -101.9117
24.3303 | -101.8654
24.334 | -101.8173
24.3376 | -101.768
24.3413 | -101.7178
24.345 | -101.6653
24.3486 | -101.6124
24.3523 | -101.558
24.3559 | -101.5019
24.3596 | -101.4454
24.3633 | -101.388
24.3669 | -101.3306
24.3706 | -101.2731
24.3742 | -101.2153
24.3779 | -101.1594
24.3815 | -101.1043
24.3852 | -101.0514
24.3889 | -101.0009
24.3925 | -100.9505
24.3962 | -100.901
24.3998 | -100.8536
24.4035 | -100.811
24.4072 | -100.772
24.4108 | -100.7344
24.4145 | -100.7003
24.4181 | -100.6699
24.4218 | -100.6439
24.4254 | -100.6213
24.4291 | -100.6014
24.4328 | -100.5848
24.4364 | -100.5728
24.4401 | -100.5651
24.4437 | -100.5622
24.4474 | -100.564
24.451 | -100.5699
24.4547 | -100.5788
24.4584 | -100.592
24.462 | -100.608
24.4657 | -100.6292
24.4693 | -100.6556
24.473 | -100.6858
24.4767 | -100.7191
24.4803 | -100.7548
24.484 | -100.7939
24.4876 | -100.8364
24.4913 | -100.8813
24.4949 | -100.9296
24.4986 | -100.9795
24.5023 | -101.0319
24.5059 | -101.0859
24.5096 | -101.143
24.5132 | -101.2032
24.5169 | -101.2635
24.5206 | -101.3254
24.5242 | -101.3881
24.5279 | -101.4513
24.5315 | -101.5148
24.5352 | -101.5791
24.5388 | -101.6437
24.5425 | -101.7074
24.5462 | -101.7703
24.5498 | -101.8339
24.5535 | -101.8952
24.5571 | -101.9553
24.5608 | -102.0125
24.5645 | -102.0673
24.5681 | -102.1166
24.5718 | -102.1594
24.5754 | -102.1977
24.5791 | -102.2279
24.5827 | -102.233
24.5864 | -102.239
24.5901 | -102.2421
24.5937 | -102.2389
24.5974 | -102.237
24.601 | -102.2378
24.6047 | -102.2408
24.6084 | -102.241
24.612 | -102.237
24.6157 | -102.2317
24.6193 | -102.2256
24.623 | -102.2183
24.6266 | -102.1993
24.6303 | -102.1718
24.634 | -102.1363
24.6376 | -102.0951
24.6413 | -102.045
24.6449 | -101.9879
24.6486 | -101.9278
24.6522 | -101.8649
24.6559 | -101.8014
24.6596 | -101.7363
24.6632 | -101.6681
24.6669 | -101.5998
24.6705 | -101.5315
24.6742 | -101.4617
24.6779 | -101.3911
24.6815 | -101.3217
24.6852 | -101.2535
24.6888 | -101.1844
24.6925 | -101.1174
24.6961 | -101.0512
24.6998 | -100.9869
24.7035 | -100.9252
24.7071 | -100.8659
24.7108 | -100.8095
24.7144 | -100.7569
24.7181 | -100.7066
24.7218 | -100.66
24.7254 | -100.6166
24.7291 | -100.5777
24.7327 | -100.5426
24.7364 | -100.5106
24.74 | -100.4831
24.7437 | -100.461
24.7474 | -100.4447
24.751 | -100.4324
24.7547 | -100.4242
24.7583 | -100.4211
24.762 | -100.4227
24.7657 | -100.4307
24.7693 | -100.443
24.773 | -100.4599
24.7766 | -100.4828
24.7803 | -100.5105
24.7839 | -100.5421
24.7876 | -100.5768
24.7913 | -100.6162
24.7949 | -100.6601
24.7986 | -100.7093
24.8022 | -100.79
24.8059 | -100.8446
24.8095 | -100.9024
24.8132 | -100.9637
24.8169 | -101.0266
24.8205 | -101.0917
24.8242 | -101.1605
24.8278 | -101.2321
24.8315 | -101.306
24.8352 | -101.3798
24.8388 | -101.4546
24.8425 | -101.5324
24.8461 | -101.6105
24.8498 | -101.6906
24.8534 | -101.77
24.8571 | -101.8502
24.8608 | -101.9291
24.8644 | -102.0069
24.8681 | -102.0847
24.8717 | -102.1608
24.8754 | -102.2361
24.8791 | -102.309
24.8827 | -102.3804
24.8864 | -102.4491
24.89 | -102.5148
24.8937 | -102.5765
24.8973 | -102.6368
24.901 | -102.6959
24.9047 | -102.7528
24.9083 | -102.8081
24.912 | -102.8614
24.9156 | -102.9123
24.9193 | -102.9644
24.923 | -103.0173
24.9266 | -103.069
24.9303 | -103.1222
24.9339 | -103.1927
24.9376 | -103.2492
24.9412 | -103.2985
24.9449 | -103.345
24.9486 | -103.3772
24.9522 | -103.3982
24.9559 | -103.4099
24.9595 | -103.4208
24.9632 | -103.4316
24.9668 | -103.4433
24.9705 | -103.4571
24.9742 | -103.4694
24.9778 | -103.4811
24.9815 | -103.4937
24.9851 | -103.5056
24.9888 | -103.5154
24.9925 | -103.523
24.9961 | -103.5299
24.9998 | -103.5385
25.0034 | -103.5488
25.0071 | -103.5576
25.0107 | -103.5672
25.0144 | -103.5777
25.0181 | -103.586
25.0217 | -103.5931
25.0254 | -103.6014
25.029 | -103.6122
25.0327 | -103.6219
25.0364 | -103.6313
25.04 | -103.6424
25.0437 | -103.6505
25.0473 | -103.6602
25.051 | -103.6688
25.0546 | -103.6748
25.0583 | -103.6816
25.062 | -103.6884
25.0656 | -103.6936
25.0693 | -103.7001
25.0729 | -103.7053
25.0766 | -103.7103
25.0803 | -103.7153
25.0839 | -103.7194
25.0876 | -103.7235
25.0912 | -103.727
25.0949 | -103.7299
25.0985 | -103.7318
25.1022 | -103.7324
25.1059 | -103.733
25.1095 | -103.7322
25.1132 | -103.7307
25.1168 | -103.729
25.1205 | -103.7288
25.1241 | -103.727
25.1278 | -103.7228
25.1315 | -103.7191
25.1351 | -103.7153
25.1388 | -103.7123
25.1424 | -103.7078
25.1461 | -103.7019
25.1498 | -103.6955
25.1534 | -103.6885
25.1571 | -103.6795
25.1607 | -103.671
25.1644 | -103.6634
25.168 | -103.6561
25.1717 | -103.6499
25.1754 | -103.6431
25.179 | -103.6341
25.1827 | -103.6245
25.1863 | -103.6163
25.19 | -103.6096
25.1937 | -103.6032
25.1973 | -103.5962
25.201 | -103.5882
25.2046 | -103.5788
25.2083 | -103.5701
25.2119 | -103.5616
25.2156 | -103.5536
25.2193 | -103.5456
25.2229 | -103.5374
25.2266 | -103.529
25.2302 | -103.5202
25.2339 | -103.5129
25.2376 | -103.5067
25.2412 | -103.5015
25.2449 | -103.4976
25.2485 | -103.4945
25.2522 | -103.4905
25.2558 | -103.4844
25.2595 | -103.4783
25.2632 | -103.4723
25.2668 | -103.4671
25.2705 | -103.4642
25.2741 | -103.4605
25.2778 | -103.4572
25.2814 | -103.4532
25.2851 | -103.4496
25.2888 | -103.4473
25.2924 | -103.4452
25.2961 | -103.442
25.2997 | -103.4369
25.3034 | -103.4327
25.3071 | -103.4298
25.3107 | -103.4279
25.3144 | -103.4279
25.318 | -103.4281
25.3217 | -103.4277
25.3253 | -103.4265
25.329 | -103.427
25.3327 | -103.4259
25.3363 | -103.4247
25.34 | -103.4237
25.3436 | -103.4224
25.3473 | -103.4233
25.351 | -103.4252
25.3546 | -103.4265
25.3583 | -103.4269
25.3619 | -103.4281
25.3656 | -103.4301
25.3692 | -103.4325
25.3729 | -103.4352
25.3766 | -103.4379
25.3802 | -103.441
25.3839 | -103.4433
25.3875 | -103.4456
25.3912 | -103.4489
25.3949 | -103.451
25.3985 | -103.4519
25.4022 | -103.4544
25.4058 | -103.4558
25.4095 | -103.4581
25.4131 | -103.4604
25.4168 | -103.4616
25.4205 | -103.4622
25.4241 | -103.4633
25.4278 | -103.4647
25.4314 | -103.4662
25.4351 | -103.4666
25.4387 | -103.4669
25.4424 | -103.4689
25.4461 | -103.4705
25.4497 | -103.4712
25.4534 | -103.472
25.457 | -103.4713
25.4607 | -103.4705
25.4644 | -103.4678
25.468 | -103.4663
25.4717 | -103.464
25.4753 | -103.4608
25.479 | -103.4576
25.4826 | -103.4556
25.4863 | -103.4549
25.49 | -103.4536
25.4936 | -103.4524
25.4973 | -103.4502
25.5009 | -103.4481
25.5046 | -103.4442
25.5083 | -103.4396
25.5119 | -103.4352
25.5156 | -103.4318
25.5192 | -103.4286
25.5229 | -103.4242
25.5265 | -103.4193
25.5302 | -103.4135
25.5339 | -103.4073
25.5375 | -103.4018
25.5412 | -103.3968
25.5448 | -103.3918
25.5485 | -103.3869
25.5522 | -103.3813
25.5558 | -103.3732
25.5595 | -103.3657
25.5631 | -103.3609
25.5668 | -103.3569
25.5704 | -103.3521
25.5741 | -103.347
25.5778 | -103.3431
25.5814 | -103.3412
25.5851 | -103.3403
25.5887 | -103.339
25.5924 | -103.3381
25.596 | -103.3369
25.5997 | -103.3345
25.6034 | -103.3323
25.607 | -103.3316
25.6107 | -103.3334
25.6143 | -103.3367
25.618 | -103.3387
25.6217 | -103.3413
25.6253 | -103.3452
25.629 | -103.348
25.6326 | -103.351
25.6363 | -103.3564
25.6399 | -103.3626
25.6436 | -103.3691
25.6473 | -103.3749
25.6509 | -103.3823
25.6546 | -103.3904
25.6582 | -103.3994
25.6619 | -103.4089
25.6656 | -103.4185
25.6692 | -103.4283
25.6729 | -103.4365
25.6765 | -103.4462
25.6802 | -103.4567
25.6838 | -103.4661
25.6875 | -103.475
25.6912 | -103.4837
25.6948 | -103.4926
25.6985 | -103.5015
25.7021 | -103.5089
25.7058 | -103.5185
25.7095 | -103.5281
25.7131 | -103.5367
25.7168 | -103.5459
25.7204 | -103.5548
25.7241 | -103.5605
25.7277 | -103.5631
25.7314 | -103.5674
25.7351 | -103.5721
25.7387 | -103.5767
25.7424 | -103.5803
25.746 | -103.5852
25.7497 | -103.5896
25.7533 | -103.5928
25.757 | -103.5974
25.7607 | -103.6006
25.7643 | -103.6029
25.768 | -103.605
25.7716 | -103.6073
25.7753 | -103.6092
25.779 | -103.6102
25.7826 | -103.6108
25.7863 | -103.6101
25.7899 | -103.6077
25.7936 | -103.6041
25.7972 | -103.6019
25.8009 | -103.5999
25.8046 | -103.5955
25.8082 | -103.5903
25.8119 | -103.5872
25.8155 | -103.5814
25.8192 | -103.5748
25.8229 | -103.57
25.8265 | -103.5679
25.8302 | -103.5651
25.8338 | -103.5614
25.8375 | -103.5573
25.8411 | -103.5507
25.8448 | -103.5435
25.8485 | -103.5365
25.8521 | -103.5304
25.8558 | -103.5255
25.8594 | -103.5195
25.8631 | -103.5145
25.8668 | -103.5083
25.8704 | -103.5044
25.8741 | -103.5011
25.8777 | -103.4966
25.8814 | -103.4928
25.885 | -103.4885
25.8887 | -103.4858
25.8924 | -103.4832
25.896 | -103.4815
25.8997 | -103.4799
25.9033 | -103.4767
25.907 | -103.473
25.9107 | -103.4702
25.9143 | -103.4682
25.918 | -103.4679
25.9216 | -103.468
25.9253 | -103.4667
25.9289 | -103.4654
25.9326 | -103.4634
25.9363 | -103.4607
25.9399 | -103.4575
25.9436 | -103.455
25.9472 | -103.4523
25.9509 | -103.4505
25.9545 | -103.4487
25.9582 | -103.4468
25.9619 | -103.4453
25.9655 | -103.4448
25.9692 | -103.4458
25.9728 | -103.4453
25.9765 | -103.4423
25.9802 | -103.4399
25.9838 | -103.4387
25.9875 | -103.4364
25.9911 | -103.4352
25.9948 | -103.4344
25.9984 | -103.4322
26.0021 | -103.4294
26.0058 | -103.4271
26.0094 | -103.424
26.0131 | -103.4226
26.0167 | -103.4212
26.0204 | -103.4192
26.0241 | -103.4174
26.0277 | -103.4139
26.0314 | -103.4103
26.035 | -103.4063
26.0387 | -103.4037
26.0423 | -103.4014
26.046 | -103.3981
26.0497 | -103.3969
26.0533 | -103.3956
26.057 | -103.3934
26.0606 | -103.3909
26.0643 | -103.3877
26.0679 | -103.3868
26.0716 | -103.3845
26.0753 | -103.3798
26.0789 | -103.3758
26.0826 | -103.3726
26.0862 | -103.3704
26.0899 | -103.3684
26.0936 | -103.37
26.0972 | -103.3684
26.1009 | -103.3672
26.1045 | -103.366
26.1082 | -103.3656
26.1118 | -103.3631
26.1155 | -103.361
26.1192 | -103.3592
26.1228 | -103.3568
26.1265 | -103.3571
26.1301 | -103.356
26.1338 | -103.3549
26.1375 | -103.3544
26.1411 | -103.3538
26.1448 | -103.3525
26.1484 | -103.3524
26.1521 | -103.3518
26.1557 | -103.3507
26.1594 | -103.3515
26.1631 | -103.3509
26.1667 | -103.3488
26.1704 | -103.3477
26.174 | -103.3471
26.1777 | -103.3457
26.1814 | -103.3474
26.185 | -103.3483
26.1887 | -103.3444
26.1923 | -103.3421
26.196 | -103.3411
26.1996 | -103.3415
26.2033 | -103.3409
26.207 | -103.3409
26.2106 | -103.3412
26.2143 | -103.3415
26.2179 | -103.3406
26.2216 | -103.3394
26.2252 | -103.3397
26.2289 | -103.3411
26.2326 | -103.3403
26.2362 | -103.3384
26.2399 | -103.3359
26.2435 | -103.3333
26.2472 | -103.3309
26.2509 | -103.3293
26.2545 | -103.3279
26.2582 | -103.3256
26.2618 | -103.3231
26.2655 | -103.3187
26.2691 | -103.3148
26.2728 | -103.3114
26.2765 | -103.3076
26.2801 | -103.3035
26.2838 | -103.2992
26.2874 | -103.293
26.2911 | -103.2881
26.2948 | -103.283
26.2984 | -103.2761
26.3021 | -103.2698
26.3057 | -103.2642
26.3094 | -103.2584
26.313 | -103.2531
26.3167 | -103.2476
26.3204 | -103.2401
26.324 | -103.2335
26.3277 | -103.2257
26.3313 | -103.2169
26.335 | -103.2087
26.3387 | -103.1997
26.3423 | -103.192
26.346 | -103.1852
26.3496 | -103.1765
26.3533 | -103.1674
26.3569 | -103.1576
26.3606 | -103.1465
26.3643 | -103.1368
26.3679 | -103.1274
26.3716 | -103.1163
26.3752 | -103.1036
26.3789 | -103.0914
26.3826 | -103.08
26.3862 | -103.0683
26.3899 | -103.0566
26.3935 | -103.0452
26.3972 | -103.0339
26.4008 | -103.0223
26.4045 | -103.0115
26.4082 | -103.0029
26.4118 | -102.9921
26.4155 | -102.9788
26.4191 | -102.9655
26.4228 | -102.9515
26.4264 | -102.9394
26.4301 | -102.9283
26.4338 | -102.9155
26.4374 | -102.9035
26.4411 | -102.8917
26.4447 | -102.88
26.4484 | -102.8691
26.4521 | -102.857
26.4557 | -102.8464
26.4594 | -102.8347
26.463 | -102.8233
26.4667 | -102.8094
26.4703 | -102.7973
26.474 | -102.7847
26.4777 | -102.7728
26.4813 | -102.7607
26.485 | -102.749
26.4886 | -102.7373
26.4923 | -102.7261
26.496 | -102.7147
26.4996 | -102.7016
26.5033 | -102.6905
26.5069 | -102.6804
26.5106 | -102.6703
26.5142 | -102.661
26.5179 | -102.6515
26.5216 | -102.6404
26.5252 | -102.629
26.5289 | -102.6173
26.5325 | -102.6066
26.5362 | -102.5969
26.5399 | -102.5882
26.5435 | -102.5785
26.5472 | -102.5699
26.5508 | -102.5609
26.5545 | -102.5512
26.5581 | -102.5422
26.5618 | -102.5356
26.5655 | -102.5285
26.5691 | -102.5209
26.5728 | -102.5126
26.5764 | -102.5068
26.5801 | -102.5023
26.5837 | -102.4966
26.5874 | -102.4886
26.5911 | -102.4817
26.5947 | -102.4761
26.5984 | -102.471
26.602 | -102.4655
26.6057 | -102.4615
26.6094 | -102.4587
26.613 | -102.4561
26.6167 | -102.4537
26.6203 | -102.4538
26.624 | -102.4524
26.6276 | -102.4512
26.6313 | -102.4508
26.635 | -102.4496
26.6386 | -102.4505
26.6423 | -102.4511
26.6459 | -102.4516
26.6496 | -102.4543
26.6533 | -102.4563
26.6569 | -102.4575
26.6606 | -102.4603
26.6642 | -102.4633
26.6679 | -102.4665
26.6715 | -102.4693
26.6752 | -102.4718
26.6789 | -102.4745
26.6825 | -102.4783
26.6862 | -102.4808
26.6898 | -102.4825
26.6935 | -102.485
26.6972 | -102.488
26.7008 | -102.4921
26.7045 | -102.4966
26.7081 | -102.5012
26.7118 | -102.5051
26.7154 | -102.5088
26.7191 | -102.5127
26.7228 | -102.5188
26.7264 | -102.5255
26.7301 | -102.5304
26.7337 | -102.5342
26.7374 | -102.5369
26.741 | -102.5421
26.7447 | -102.5473
26.7484 | -102.5522
26.752 | -102.5563
26.7557 | -102.5574
26.7593 | -102.5583
26.763 | -102.5601
26.7667 | -102.5608
26.7703 | -102.5632
26.774 | -102.566
26.7776 | -102.5688
26.7813 | -102.5702
26.7849 | -102.5694
26.7886 | -102.5702
26.7923 | -102.5724
26.7959 | -102.5755
26.7996 | -102.5775
26.8032 | -102.5762
26.8069 | -102.5759
26.8106 | -102.5776
26.8142 | -102.5784
26.8179 | -102.5776
26.8215 | -102.5751
26.8252 | -102.5729
26.8288 | -102.5733
26.8325 | -102.573
26.8362 | -102.5732
26.8398 | -102.5729
26.8435 | -102.574
26.8471 | -102.5742
26.8508 | -102.5755
26.8545 | -102.5756
26.8581 | -102.5758
26.8618 | -102.576
26.8654 | -102.5755
26.8691 | -102.5752
26.8727 | -102.5767
26.8764 | -102.5782
26.8801 | -102.5806
26.8837 | -102.5817
26.8874 | -102.5827
26.891 | -102.5848
26.8947 | -102.588
26.8983 | -102.5923
26.902 | -102.5987
26.9057 | -102.6053
26.9093 | -102.6099
26.913 | -102.6167
26.9166 | -102.6216
26.9203 | -102.6277
26.924 | -102.6342
26.9276 | -102.6394
26.9313 | -102.6464
26.9349 | -102.6537
26.9386 | -102.6624
26.9422 | -102.6717
26.9459 | -102.6789
26.9496 | -102.6864
26.9532 | -102.6952
26.9569 | -102.7025
26.9605 | -102.7099
26.9642 | -102.7204
26.9679 | -102.7332
26.9715 | -102.745
26.9752 | -102.7557
26.9788 | -102.766
26.9825 | -102.7772
26.9861 | -102.7868
26.9898 | -102.7976
26.9935 | -102.8074
26.9971 | -102.8175
27.0008 | -102.8256
27.0044 | -102.8342
27.0081 | -102.8439
27.0118 | -102.8528
27.0154 | -102.862
27.0191 | -102.8694
27.0227 | -102.877
27.0264 | -102.8854
27.03 | -102.8921
27.0337 | -102.8989
27.0374 | -102.9059
27.041 | -102.9102
27.0447 | -102.9146
27.0483 | -102.9196
27.052 | -102.9248
27.0556 | -102.9298
27.0593 | -102.9336
27.063 | -102.9358
27.0666 | -102.9375
27.0703 | -102.9399
27.0739 | -102.9428
27.0776 | -102.9445
27.0813 | -102.9443
27.0849 | -102.9437
27.0886 | -102.9422
27.0922 | -102.9406
27.0959 | -102.9392
27.0995 | -102.9359
27.1032 | -102.9311
27.1069 | -102.9259
27.1105 | -102.9218
27.1142 | -102.9162
27.1178 | -102.9115
27.1215 | -102.9039
27.1252 | -102.897
27.1288 | -102.892
27.1325 | -102.8856
27.1361 | -102.8781
27.1398 | -102.8698
27.1434 | -102.86
27.1471 | -102.8518
27.1508 | -102.8414
27.1544 | -102.8301
27.1581 | -102.8195
27.1617 | -102.8088
27.1654 | -102.7991
27.1691 | -102.7895
27.1727 | -102.7794
27.1764 | -102.7693
27.18 | -102.7571
27.1837 | -102.7456
27.1873 | -102.7333
27.191 | -102.7222
27.1947 | -102.711
27.1983 | -102.6999
27.202 | -102.6903
27.2056 | -102.6807
27.2093 | -102.6687
27.2129 | -102.6553
27.2166 | -102.642
27.2203 | -102.6317
27.2239 | -102.6201
27.2276 | -102.6076
27.2312 | -102.5954
27.2349 | -102.5834
27.2386 | -102.5728
27.2422 | -102.5621
27.2459 | -102.5522
27.2495 | -102.5425
27.2532 | -102.5319
27.2568 | -102.5224
27.2605 | -102.5126
27.2642 | -102.5033
27.2678 | -102.4931
27.2715 | -102.4834
27.2751 | -102.4754
27.2788 | -102.467
27.2825 | -102.4569
27.2861 | -102.4488
27.2898 | -102.4414
27.2934 | -102.4347
27.2971 | -102.428
27.3007 | -102.4198
27.3044 | -102.4119
27.3081 | -102.4058
27.3117 | -102.3997
27.3154 | -102.3927
27.319 | -102.3865
27.3227 | -102.3817
27.3264 | -102.375
27.33 | -102.3702
27.3337 | -102.366
27.3373 | -102.3618
27.341 | -102.3579
27.3446 | -102.3534
27.3483 | -102.3484
27.352 | -102.3429
27.3556 | -102.3382
27.3593 | -102.3346
27.3629 | -102.3316
27.3666 | -102.3277
27.3702 | -102.3248
27.3739 | -102.3207
27.3776 | -102.3165
27.3812 | -102.3133
27.3849 | -102.31
27.3885 | -102.3065
27.3922 | -102.3016
27.3959 | -102.2953
27.3995 | -102.291
27.4032 | -102.2878
27.4068 | -102.284
27.4105 | -102.2787
27.4141 | -102.2731
27.4178 | -102.2666
27.4215 | -102.2619
27.4251 | -102.2586
27.4288 | -102.2545
27.4324 | -102.2498
27.4361 | -102.245
27.4398 | -102.2391
27.4434 | -102.2331
27.4471 | -102.2288
27.4507 | -102.2236
27.4544 | -102.2179
27.458 | -102.213
27.4617 | -102.207
27.4654 | -102.2009
27.469 | -102.1969
27.4727 | -102.1934
27.4763 | -102.1885
27.48 | -102.1828
27.4837 | -102.1776
27.4873 | -102.1735
27.491 | -102.1707
27.4946 | -102.1673
27.4983 | -102.1644
27.5019 | -102.1613
27.5056 | -102.1566
27.5093 | -102.1525
27.5129 | -102.1479
27.5166 | -102.1421
27.5202 | -102.1388
27.5239 | -102.1371
27.5276 | -102.1333
27.5312 | -102.1276
27.5349 | -102.1221
27.5385 | -102.1176
27.5422 | -102.1132
27.5458 | -102.1077
27.5495 | -102.1022
27.5532 | -102.0973
27.5568 | -102.0928
27.5605 | -102.0878
27.5641 | -102.0822
27.5678 | -102.0798
27.5714 | -102.0763
27.5751 | -102.0702
27.5788 | -102.0636
27.5824 | -102.0573
27.5861 | -102.0521
27.5897 | -102.0483
27.5934 | -102.0449
27.5971 | -102.0425
27.6007 | -102.0396
27.6044 | -102.0377
27.608 | -102.0336
27.6117 | -102.0297
27.6153 | -102.0254
27.619 | -102.0197
27.6227 | -102.0155
27.6263 | -102.0115
27.63 | -102.0077
27.6336 | -102.0041
27.6373 | -101.9999
27.641 | -101.9964
27.6446 | -101.9932
27.6483 | -101.9891
27.6519 | -101.9858
27.6556 | -101.9832
27.6592 | -101.9809
27.6629 | -101.9785
27.6666 | -101.9757
27.6702 | -101.9734
27.6739 | -101.9714
27.6775 | -101.97
27.6812 | -101.9689
27.6849 | -101.968
27.6885 | -101.9664
27.6922 | -101.9649
27.6958 | -101.9642
27.6995 | -101.9654
27.7031 | -101.9665
27.7068 | -101.9671
27.7105 | -101.9674
27.7141 | -101.9688
27.7178 | -101.9697
27.7214 | -101.9721
27.7251 | -101.9749
27.7287 | -101.9774
27.7324 | -101.9818
27.7361 | -101.9869
27.7397 | -101.9921
27.7434 | -101.9963
27.747 | -102
27.7507 | -102.005
27.7544 | -102.0107
27.758 | -102.0165
27.7617 | -102.0235
27.7653 | -102.0302
27.769 | -102.0366
27.7726 | -102.0445
27.7763 | -102.0528
27.78 | -102.0623
27.7836 | -102.073
27.7873 | -102.0813
27.7909 | -102.0891
27.7946 | -102.0981
27.7983 | -102.1077
27.8019 | -102.1171
27.8056 | -102.1266
27.8092 | -102.1349
27.8129 | -102.1429
27.8165 | -102.1506
27.8202 | -102.1598
27.8239 | -102.1689
27.8275 | -102.1788
27.8312 | -102.19
27.8348 | -102.2001
27.8385 | -102.2082
27.8422 | -102.2166
27.8458 | -102.2252
27.8495 | -102.2324
27.8531 | -102.2391
27.8568 | -102.2435
27.8604 | -102.2462
27.8641 | -102.2498
27.8678 | -102.2562
27.8714 | -102.2637
27.8751 | -102.269
27.8787 | -102.2739
27.8824 | -102.2778
27.886 | -102.282
27.8897 | -102.2863
27.8934 | -102.291
27.897 | -102.296
27.9007 | -102.3007
27.9043 | -102.3046
27.908 | -102.3077
27.9117 | -102.3104
27.9153 | -102.3114
27.919 | -102.3138
27.9226 | -102.3178
27.9263 | -102.3174
27.9299 | -102.3101
27.9336 | -102.3052
27.9373 | -102.3046
27.9409 | -102.3056
27.9446 | -102.3054
27.9482 | -102.3033
27.9519 | -102.3017
27.9556 | -102.3013
27.9592 | -102.2997
27.9629 | -102.3
27.9665 | -102.3006
27.9702 | -102.299
27.9738 | -102.298
27.9775 | -102.2968
27.9812 | -102.2953
27.9848 | -102.2937
27.9885 | -102.2911
27.9921 | -102.287
27.9958 | -102.2836
27.9995 | -102.28
28.0031 | -102.2769