## VERIFICATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 270 |  | CTBTO Vienna International Centre Wagramestr.,5 A-1400 VIENNA Austria
Transducer Block | 02-90C/P | 277
Hood | 02-92B/P | 217
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE's REFERENCE
Atmospheric pressure | 1009.8 | hPa |  | PRC PRO 01
Humidity | 53.4 | %/Hr
Internal temperature of calibration bench | 24.34 | °C |  | PLACES OF INTERVENTION
Auxiliary sensor pressure | 1019 | hPa |  | SEISMO WAVE, 22300 LANNION
Auxiliary sensor temperature | 23.7 | °C
ADJUSTMENT OF SENSITIVITIES AT 1 Hz
Output/Input Parameter | Theoretical Value | Measured Value @1Hz | Measurement Uncertainty (%) | Pass/Fail criteria Margin (%) | Unit | Results
Sensitivity of Pressure Derivative output | 2 | 2 | +/- 4 | +/- 5 | mV/Pa/sec | OK
Sensitivity of Pressure output | 20 | 19.94 | +/- 4 | +/- 5 | mV/Pa | OK
Sensitivity of Calibration coil | 6 | 5.97 | +/- 4 | +/- 5 | Pa/V | OK
Input Parameter | Average Value | Measured Value | Measurement Uncertainty (%) |  | Unit
Impedance of calibration input | 3100 | 3110 | +/- 0,3 |  | Ohm
MATERIAL USED
Type of operation | Instrument |  |  | Type
Adjustment | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Metrix Multimeter |  |  | 34410A
| Oscilloscope |  |  | Tektronix TDS 3014B
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the acoustical transfer function | Microphone |  |  | 4193-L-004
| Microphone power supply |  |  | 2829
| Meteo station |  |  | ALMEMO 2690
| Power Supply |  |  | E36104A
| Acoustical and electrical measurement system |  |  | PR 000488 -
| Acoustical generator |  |  | TMG_ASM_3-22a
Measurement of the electrical transfer function | Digitizer |  |  | Digitizer Hood
| Power Supply |  |  | E36104A
Self-noise measurement | Digitizer |  |  | TAURUS
| Power Supply |  |  | E36104A
| Vibration isolation |  |  | Accurion Nano 30
PERSONS IN CHARGE OF
Qualification | Name | Operation |  | Date | VISA
Technician | Ludovic LECLAIR | Adjustment |  | 18/07/2018
|  | Measurement of electrical transfer function Measurement of acoustical transfer function Measurement of self-noise |  | 20/07/2018
Technical Manager | Anthony HUE | Approval |  | 24/07/2018
CEO | Robert GLEMOT | Approval |  | 24/07/2018
DECISION OF CONFORMITY
This sensor is compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure |  |  |  | X
This sensor is not compliant with the technical requirements defined into the "PRC PRO 01" manufacturer procedure
This document contains 4 pages. These pages are filled following the type of equipement and according the services required by clients
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which defined the verification report. It can be used to demonstrate the link between the measurement methods to the standard national or international references, provided it meets the recommendations of the documentation booklet X 07-015

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.0061043 | -39.5207 | 57.599
0.012209 | -36.0831 | 38.4637
0.018313 | -35.0142 | 28.3034
0.024417 | -34.5489 | 22.037
0.030521 | -34.4328 | 18.2499
0.036626 | -34.2873 | 15.186
0.04273 | -34.192 | 12.8162
0.048834 | -34.1457 | 11.2778
0.054938 | -34.0808 | 10.1344
0.061043 | -34.021 | 9.0535
0.067147 | -34.0108 | 8.2322
0.073251 | -34.0113 | 7.4726
0.079355 | -34.0285 | 6.8099
0.08546 | -34.0414 | 6.214
0.091564 | -34.0364 | 5.7692
0.097668 | -34.0188 | 5.5133
0.10377 | -33.9975 | 5.262
0.10988 | -33.9808 | 4.9819
0.11598 | -33.9719 | 4.6876
0.12209 | -33.9736 | 4.4027
0.12819 | -33.9833 | 4.1191
0.13429 | -33.9941 | 3.8326
0.1404 | -33.9977 | 3.557
0.1465 | -34 | 3.301
0.15261 | -33.9977 | 3.0816
0.15871 | -33.9917 | 2.9184
0.16482 | -33.984 | 2.813
0.17092 | -33.9817 | 2.7699
0.17702 | -33.9795 | 2.6951
0.18313 | -33.9786 | 2.615
0.18923 | -33.9781 | 2.5059
0.19534 | -33.9769 | 2.3798
0.20144 | -33.9757 | 2.239
0.20754 | -33.9729 | 2.1186
0.21365 | -33.972 | 2.0043
0.21975 | -33.9716 | 1.9061
0.22586 | -33.9726 | 1.8185
0.23196 | -33.975 | 1.7456
0.23807 | -33.979 | 1.6826
0.24417 | -33.983 | 1.6248
0.25027 | -33.987 | 1.5631
0.25638 | -33.9894 | 1.5021
0.26248 | -33.9913 | 1.4367
0.26859 | -33.9923 | 1.3726
0.27469 | -33.9927 | 1.3027
0.2808 | -33.9938 | 1.2298
0.2869 | -33.9946 | 1.1596
0.293 | -33.9956 | 1.093
0.29911 | -33.9965 | 1.0329
0.30521 | -33.9975 | 0.97359
0.31132 | -33.9983 | 0.93359
0.31742 | -33.9996 | 0.88967
0.32353 | -34 | 0.8533
0.32963 | -34.0009 | 0.82261
0.33573 | -34.0019 | 0.80034
0.34184 | -34.0033 | 0.77729
0.34794 | -34.0054 | 0.75228
0.35405 | -34.0074 | 0.72685
0.36015 | -34.0099 | 0.70034
0.36626 | -34.0118 | 0.67202
0.37236 | -34.0139 | 0.63492
0.37846 | -34.0159 | 0.59061
0.38457 | -34.0168 | 0.53421
0.39067 | -34.0178 | 0.48774
0.39678 | -34.0185 | 0.43633
0.40288 | -34.019 | 0.38285
0.40899 | -34.0189 | 0.33157
0.41509 | -34.0183 | 0.28123
0.42119 | -34.0184 | 0.23832
0.4273 | -34.0184 | 0.19258
0.4334 | -34.0184 | 0.14692
0.43951 | -34.0188 | 0.10394
0.44561 | -34.0197 | 0.064429
0.45172 | -34.0212 | 0.032735
0.45782 | -34.0221 | -0.0041375
0.46392 | -34.0233 | -0.040135
0.47003 | -34.0244 | -0.072433
0.47613 | -34.0252 | -0.10646
0.48224 | -34.0257 | -0.14112
0.48834 | -34.0248 | -0.17575
0.49445 | -34.0246 | -0.21085
0.50055 | -34.024 | -0.24295
0.50665 | -34.0233 | -0.277
0.51276 | -34.0223 | -0.31092
0.51886 | -34.0213 | -0.34468
0.52497 | -34.0202 | -0.36899
0.53107 | -34.0194 | -0.39843
0.53717 | -34.0187 | -0.42658
0.54328 | -34.0179 | -0.44959
0.54938 | -34.0178 | -0.47219
0.55549 | -34.0178 | -0.49335
0.56159 | -34.0189 | -0.50554
0.5677 | -34.0192 | -0.5206
0.5738 | -34.0203 | -0.53735
0.5799 | -34.0211 | -0.55276
0.58601 | -34.0221 | -0.56356
0.59211 | -34.0238 | -0.5708
0.59822 | -34.0251 | -0.58107
0.60432 | -34.0266 | -0.59382
0.61043 | -34.0279 | -0.6026
0.61653 | -34.0291 | -0.61365
0.62263 | -34.0307 | -0.62379
0.62874 | -34.0314 | -0.63507
0.63484 | -34.0322 | -0.64506
0.64095 | -34.0328 | -0.65592
0.64705 | -34.0332 | -0.66635
0.65316 | -34.0334 | -0.67936
0.65926 | -34.0336 | -0.69562
0.66536 | -34.0326 | -0.71329
0.67147 | -34.0322 | -0.72691
0.67757 | -34.0319 | -0.74163
0.68368 | -34.0313 | -0.76072
0.68978 | -34.0306 | -0.77807
0.69589 | -34.0297 | -0.79723
0.70199 | -34.0288 | -0.81997
0.70809 | -34.0281 | -0.83926
0.7142 | -34.0274 | -0.85982
0.7203 | -34.0264 | -0.87912
0.72641 | -34.0258 | -0.90164
0.73251 | -34.0252 | -0.92176
0.73862 | -34.0255 | -0.9436
0.74472 | -34.0252 | -0.96375
0.75082 | -34.025 | -0.98595
0.75693 | -34.0254 | -1.0079
0.76303 | -34.0259 | -1.0299
0.76914 | -34.0277 | -1.0505
0.77524 | -34.0287 | -1.0703
0.78135 | -34.03 | -1.0905
0.78745 | -34.0313 | -1.1105
0.79355 | -34.0326 | -1.1304
0.79966 | -34.0343 | -1.15
0.80576 | -34.0365 | -1.1697
0.81187 | -34.0381 | -1.1912
0.81797 | -34.04 | -1.2125
0.82408 | -34.0417 | -1.2331
0.83018 | -34.0434 | -1.2536
0.83628 | -34.0452 | -1.2721
0.84239 | -34.0464 | -1.2925
0.84849 | -34.0478 | -1.314
0.8546 | -34.0492 | -1.3337
0.8607 | -34.0505 | -1.3544
0.86681 | -34.0515 | -1.3733
0.87291 | -34.0525 | -1.392
0.87901 | -34.053 | -1.4099
0.88512 | -34.0537 | -1.428
0.89122 | -34.0544 | -1.4463
0.89733 | -34.0548 | -1.4619
0.90343 | -34.0552 | -1.4795
0.90953 | -34.0554 | -1.4939
0.91564 | -34.0554 | -1.5098
0.92174 | -34.0557 | -1.5249
0.92785 | -34.0558 | -1.5395
0.93395 | -34.0558 | -1.5535
0.94006 | -34.056 | -1.5665
0.94616 | -34.0558 | -1.5786
0.95226 | -34.0559 | -1.5914
0.95837 | -34.056 | -1.6054
0.96447 | -34.0562 | -1.6184
0.97058 | -34.056 | -1.632
0.97668 | -34.0558 | -1.6455
0.98279 | -34.0556 | -1.6584
0.98889 | -34.0554 | -1.6693
0.99499 | -34.0554 | -1.6829
1.0011 | -34.0553 | -1.6938
1.0072 | -34.0549 | -1.7057
1.0133 | -34.0546 | -1.7167
1.0194 | -34.0543 | -1.7299
1.0255 | -34.054 | -1.7437
1.0316 | -34.0535 | -1.7559
1.0377 | -34.0531 | -1.7698
1.0438 | -34.0525 | -1.7826
1.0499 | -34.0518 | -1.798
1.056 | -34.0512 | -1.8124
1.0621 | -34.0506 | -1.8269
1.0682 | -34.0502 | -1.8395
1.0743 | -34.0496 | -1.8551
1.0805 | -34.0492 | -1.8696
1.0866 | -34.0483 | -1.8848
1.0927 | -34.0475 | -1.8995
1.0988 | -34.047 | -1.9144
1.1049 | -34.0462 | -1.9288
1.111 | -34.0455 | -1.9443
1.1171 | -34.0448 | -1.959
1.1232 | -34.0444 | -1.9743
1.1293 | -34.044 | -1.9885
1.1354 | -34.0438 | -2.0037
1.1415 | -34.0435 | -2.0193
1.1476 | -34.0429 | -2.0349
1.1537 | -34.0424 | -2.0497
1.1598 | -34.0426 | -2.0647
1.1659 | -34.0424 | -2.0782
1.172 | -34.0422 | -2.0923
1.1781 | -34.0419 | -2.1064
1.1842 | -34.0418 | -2.1207
1.1903 | -34.0422 | -2.132
1.1964 | -34.0421 | -2.1455
1.2025 | -34.0422 | -2.1596
1.2086 | -34.0425 | -2.1723
1.2147 | -34.0426 | -2.185
1.2209 | -34.0429 | -2.1965
1.227 | -34.0436 | -2.2075
1.2331 | -34.044 | -2.2196
1.2392 | -34.0444 | -2.2316
1.2453 | -34.0448 | -2.2428
1.2514 | -34.0452 | -2.253
1.2575 | -34.0458 | -2.2641
1.2636 | -34.0466 | -2.2731
1.2697 | -34.0472 | -2.2826
1.2758 | -34.0479 | -2.2938
1.2819 | -34.0486 | -2.3039
1.288 | -34.0492 | -2.3148
1.2941 | -34.05 | -2.3235
1.3002 | -34.0507 | -2.3316
1.3063 | -34.0512 | -2.3409
1.3124 | -34.0519 | -2.35
1.3185 | -34.0526 | -2.3592
1.3246 | -34.0533 | -2.3683
1.3307 | -34.0541 | -2.377
1.3368 | -34.0549 | -2.3858
1.3429 | -34.0556 | -2.3948
1.349 | -34.0561 | -2.4042
1.3551 | -34.0568 | -2.4136
1.3613 | -34.0573 | -2.422
1.3674 | -34.0579 | -2.432
1.3735 | -34.0584 | -2.4418
1.3796 | -34.0589 | -2.451
1.3857 | -34.0592 | -2.4603
1.3918 | -34.0598 | -2.4706
1.3979 | -34.0603 | -2.4798
1.404 | -34.0608 | -2.4912
1.4101 | -34.0613 | -2.5017
1.4162 | -34.0617 | -2.5123
1.4223 | -34.0622 | -2.5237
1.4284 | -34.0627 | -2.5338
1.4345 | -34.0629 | -2.5454
1.4406 | -34.0633 | -2.559
1.4467 | -34.0637 | -2.5706
1.4528 | -34.0642 | -2.5821
1.4589 | -34.0647 | -2.594
1.465 | -34.0649 | -2.605
1.4711 | -34.0655 | -2.6165
1.4772 | -34.066 | -2.6298
1.4833 | -34.0665 | -2.6429
1.4894 | -34.067 | -2.6552
1.4955 | -34.0675 | -2.6672
1.5016 | -34.0681 | -2.6798
1.5078 | -34.0685 | -2.6938
1.5139 | -34.0689 | -2.7069
1.52 | -34.0693 | -2.7197
1.5261 | -34.0697 | -2.7324
1.5322 | -34.07 | -2.7465
1.5383 | -34.0704 | -2.7592
1.5444 | -34.071 | -2.772
1.5505 | -34.0714 | -2.7855
1.5566 | -34.072 | -2.7986
1.5627 | -34.0724 | -2.8125
1.5688 | -34.073 | -2.8261
1.5749 | -34.0734 | -2.839
1.581 | -34.0737 | -2.8525
1.5871 | -34.0741 | -2.8663
1.5932 | -34.0744 | -2.8793
1.5993 | -34.0749 | -2.8937
1.6054 | -34.0752 | -2.9071
1.6115 | -34.0757 | -2.9213
1.6176 | -34.0759 | -2.934
1.6237 | -34.0762 | -2.9478
1.6298 | -34.0764 | -2.9606
1.6359 | -34.0768 | -2.9735
1.642 | -34.0772 | -2.9857
1.6482 | -34.0774 | -2.9989
1.6543 | -34.0777 | -3.0128
1.6604 | -34.0781 | -3.025
1.6665 | -34.0782 | -3.0377
1.6726 | -34.0784 | -3.0504
1.6787 | -34.0787 | -3.0632
1.6848 | -34.0786 | -3.076
1.6909 | -34.0789 | -3.0889
1.697 | -34.0791 | -3.1014
1.7031 | -34.0792 | -3.1142
1.7092 | -34.0792 | -3.1266
1.7153 | -34.0793 | -3.1388
1.7214 | -34.0792 | -3.1512
1.7275 | -34.0791 | -3.164
1.7336 | -34.0791 | -3.1771
1.7397 | -34.0792 | -3.1896
1.7458 | -34.0792 | -3.202
1.7519 | -34.0792 | -3.2144
1.758 | -34.0793 | -3.2266
1.7641 | -34.0795 | -3.2383
1.7702 | -34.0794 | -3.2515
1.7763 | -34.0794 | -3.2632
1.7824 | -34.0794 | -3.2758
1.7885 | -34.0794 | -3.2884
1.7947 | -34.0795 | -3.3006
1.8008 | -34.0795 | -3.3121
1.8069 | -34.0796 | -3.3239
1.813 | -34.0797 | -3.3363
1.8191 | -34.0796 | -3.3491
1.8252 | -34.0798 | -3.3619
1.8313 | -34.0799 | -3.3743
1.8374 | -34.0801 | -3.3863
1.8435 | -34.0801 | -3.3983
1.8496 | -34.0802 | -3.4114
1.8557 | -34.0804 | -3.4242
1.8618 | -34.0807 | -3.4365
1.8679 | -34.0808 | -3.4486
1.874 | -34.0809 | -3.4607
1.8801 | -34.0811 | -3.4731
1.8862 | -34.0814 | -3.4852
1.8923 | -34.0815 | -3.4974
1.8984 | -34.0817 | -3.5095
1.9045 | -34.0818 | -3.5212
1.9106 | -34.0819 | -3.5334
1.9167 | -34.0819 | -3.5456
1.9228 | -34.0821 | -3.5576
1.9289 | -34.0823 | -3.5694
1.9351 | -34.0824 | -3.5803
1.9412 | -34.0827 | -3.5916
1.9473 | -34.083 | -3.6037
1.9534 | -34.0832 | -3.6152
1.9595 | -34.0834 | -3.6271
1.9656 | -34.0837 | -3.6393
1.9717 | -34.084 | -3.6511
1.9778 | -34.0843 | -3.6625
1.9839 | -34.0846 | -3.674
1.99 | -34.0847 | -3.6851
1.9961 | -34.0849 | -3.6973
2.0022 | -34.0851 | -3.7087
2.0083 | -34.0852 | -3.72
2.0144 | -34.0855 | -3.7315
2.0205 | -34.0858 | -3.7432
2.0266 | -34.0862 | -3.7547
2.0327 | -34.0864 | -3.7662
2.0388 | -34.0868 | -3.7778
2.0449 | -34.0871 | -3.7892
2.051 | -34.0873 | -3.8006
2.0571 | -34.0877 | -3.812
2.0632 | -34.0879 | -3.8236
2.0693 | -34.0881 | -3.8357
2.0754 | -34.0884 | -3.8472
2.0816 | -34.0887 | -3.8593
2.0877 | -34.0889 | -3.8709
2.0938 | -34.0893 | -3.883
2.0999 | -34.0895 | -3.8947
2.106 | -34.0897 | -3.9064
2.1121 | -34.0898 | -3.9185
2.1182 | -34.0901 | -3.9293
2.1243 | -34.0904 | -3.9402
2.1304 | -34.0906 | -3.9513
2.1365 | -34.0909 | -3.9631
2.1426 | -34.0911 | -3.9752
2.1487 | -34.0914 | -3.9865
2.1548 | -34.0918 | -3.9983
2.1609 | -34.0922 | -4.0098
2.167 | -34.0926 | -4.0211
2.1731 | -34.093 | -4.0332
2.1792 | -34.0932 | -4.0443
2.1853 | -34.0936 | -4.0562
2.1914 | -34.0938 | -4.0677
2.1975 | -34.0942 | -4.0792
2.2036 | -34.0944 | -4.0914
2.2097 | -34.0947 | -4.1029
2.2158 | -34.0951 | -4.1136
2.222 | -34.0953 | -4.1241
2.2281 | -34.0956 | -4.1355
2.2342 | -34.0959 | -4.147
2.2403 | -34.0961 | -4.1577
2.2464 | -34.0963 | -4.1686
2.2525 | -34.0966 | -4.1801
2.2586 | -34.0969 | -4.1907
2.2647 | -34.0972 | -4.2013
2.2708 | -34.0975 | -4.2126
2.2769 | -34.0979 | -4.2234
2.283 | -34.098 | -4.2346
2.2891 | -34.0983 | -4.2456
2.2952 | -34.0984 | -4.2562
2.3013 | -34.0986 | -4.2672
2.3074 | -34.0989 | -4.2777
2.3135 | -34.0991 | -4.2892
2.3196 | -34.0993 | -4.2994
2.3257 | -34.0995 | -4.3101
2.3318 | -34.0997 | -4.3212
2.3379 | -34.0999 | -4.3321
2.344 | -34.1 | -4.3435
2.3501 | -34.1002 | -4.3543
2.3562 | -34.1004 | -4.3648
2.3623 | -34.1005 | -4.3753
2.3685 | -34.1006 | -4.3859
2.3746 | -34.1007 | -4.3972
2.3807 | -34.1009 | -4.4078
2.3868 | -34.1011 | -4.4186
2.3929 | -34.1013 | -4.4293
2.399 | -34.1014 | -4.4398
2.4051 | -34.1017 | -4.4504
2.4112 | -34.1018 | -4.4612
2.4173 | -34.102 | -4.4719
2.4234 | -34.1021 | -4.4827
2.4295 | -34.1022 | -4.4937
2.4356 | -34.1024 | -4.5046
2.4417 | -34.1025 | -4.5157
2.4478 | -34.1026 | -4.5267
2.4539 | -34.1028 | -4.537
2.46 | -34.1029 | -4.5485
2.4661 | -34.1031 | -4.5597
2.4722 | -34.1033 | -4.57
2.4783 | -34.1033 | -4.5805
2.4844 | -34.1034 | -4.5917
2.4905 | -34.1035 | -4.6027
2.4966 | -34.1037 | -4.6134
2.5027 | -34.1038 | -4.624
2.5089 | -34.1039 | -4.6349
2.515 | -34.104 | -4.6458
2.5211 | -34.1041 | -4.6567
2.5272 | -34.1042 | -4.6677
2.5333 | -34.1043 | -4.6787
2.5394 | -34.1045 | -4.6895
2.5455 | -34.1046 | -4.7009
2.5516 | -34.1048 | -4.7114
2.5577 | -34.105 | -4.7229
2.5638 | -34.1051 | -4.7347
2.5699 | -34.1052 | -4.7462
2.576 | -34.1052 | -4.7578
2.5821 | -34.1054 | -4.7691
2.5882 | -34.1055 | -4.7803
2.5943 | -34.1056 | -4.7913
2.6004 | -34.1058 | -4.8023
2.6065 | -34.1059 | -4.8128
2.6126 | -34.1061 | -4.8234
2.6187 | -34.1062 | -4.834
2.6248 | -34.1063 | -4.845
2.6309 | -34.1065 | -4.8565
2.637 | -34.1068 | -4.8682
2.6431 | -34.1069 | -4.8788
2.6492 | -34.107 | -4.8903
2.6554 | -34.1073 | -4.9012
2.6615 | -34.1074 | -4.9123
2.6676 | -34.1076 | -4.9238
2.6737 | -34.1079 | -4.9351
2.6798 | -34.1082 | -4.9463
2.6859 | -34.1084 | -4.9577
2.692 | -34.1086 | -4.9687
2.6981 | -34.1088 | -4.9799
2.7042 | -34.109 | -4.9912
2.7103 | -34.1092 | -5.0019
2.7164 | -34.1095 | -5.013
2.7225 | -34.1097 | -5.0245
2.7286 | -34.11 | -5.0356
2.7347 | -34.1102 | -5.0463
2.7408 | -34.1105 | -5.0565
2.7469 | -34.1107 | -5.0679
2.753 | -34.111 | -5.0786
2.7591 | -34.1113 | -5.09
2.7652 | -34.1115 | -5.1016
2.7713 | -34.1117 | -5.1125
2.7774 | -34.112 | -5.1236
2.7835 | -34.1122 | -5.1347
2.7896 | -34.1124 | -5.1454
2.7958 | -34.1126 | -5.1563
2.8019 | -34.113 | -5.1668
2.808 | -34.1132 | -5.1778
2.8141 | -34.1135 | -5.1891
2.8202 | -34.1139 | -5.2001
2.8263 | -34.1142 | -5.2111
2.8324 | -34.1145 | -5.2226
2.8385 | -34.1149 | -5.2336
2.8446 | -34.1153 | -5.2445
2.8507 | -34.1156 | -5.2554
2.8568 | -34.116 | -5.2661
2.8629 | -34.1163 | -5.2768
2.869 | -34.1166 | -5.288
2.8751 | -34.1169 | -5.2989
2.8812 | -34.1173 | -5.3103
2.8873 | -34.1176 | -5.3213
2.8934 | -34.1179 | -5.3325
2.8995 | -34.1183 | -5.3437
2.9056 | -34.1185 | -5.3549
2.9117 | -34.1189 | -5.3661
2.9178 | -34.1193 | -5.3769
2.9239 | -34.1196 | -5.3886
2.93 | -34.1201 | -5.4001
2.9361 | -34.1203 | -5.4111
2.9423 | -34.1207 | -5.4219
2.9484 | -34.1211 | -5.4331
2.9545 | -34.1215 | -5.4442
2.9606 | -34.1219 | -5.4556
2.9667 | -34.1222 | -5.4669
2.9728 | -34.1226 | -5.4782
2.9789 | -34.123 | -5.4895
2.985 | -34.1234 | -5.5008
2.9911 | -34.1238 | -5.5121
2.9972 | -34.1242 | -5.5233
3.0033 | -34.1246 | -5.5349
3.0094 | -34.125 | -5.5461
3.0155 | -34.1254 | -5.5571
3.0216 | -34.1258 | -5.5684
3.0277 | -34.1261 | -5.5802
3.0338 | -34.1265 | -5.5916
3.0399 | -34.1269 | -5.603
3.046 | -34.1272 | -5.6147
3.0521 | -34.1276 | -5.6258
3.0582 | -34.1279 | -5.6378
3.0643 | -34.1282 | -5.6493
3.0704 | -34.1285 | -5.6612
3.0765 | -34.1288 | -5.6723
3.0827 | -34.1291 | -5.6836
3.0888 | -34.1294 | -5.6949
3.0949 | -34.1298 | -5.7063
3.101 | -34.1302 | -5.7177
3.1071 | -34.1305 | -5.7295
3.1132 | -34.1309 | -5.7413
3.1193 | -34.1312 | -5.7528
3.1254 | -34.1316 | -5.7642
3.1315 | -34.132 | -5.7761
3.1376 | -34.1322 | -5.7875
3.1437 | -34.1326 | -5.7991
3.1498 | -34.1329 | -5.8111
3.1559 | -34.1332 | -5.8226
3.162 | -34.1335 | -5.8347
3.1681 | -34.1337 | -5.8466
3.1742 | -34.1341 | -5.8582
3.1803 | -34.1344 | -5.8699
3.1864 | -34.1347 | -5.8811
3.1925 | -34.135 | -5.8927
3.1986 | -34.1353 | -5.9046
3.2047 | -34.1356 | -5.9166
3.2108 | -34.1358 | -5.9277
3.2169 | -34.1361 | -5.9394
3.223 | -34.1364 | -5.9511
3.2292 | -34.1366 | -5.9629
3.2353 | -34.1368 | -5.9737
3.2414 | -34.1371 | -5.9852
3.2475 | -34.1373 | -5.9966
3.2536 | -34.1376 | -6.008
3.2597 | -34.1378 | -6.019
3.2658 | -34.138 | -6.03
3.2719 | -34.1383 | -6.041
3.278 | -34.1385 | -6.0526
3.2841 | -34.1388 | -6.0637
3.2902 | -34.139 | -6.0752
3.2963 | -34.1393 | -6.0866
3.3024 | -34.1394 | -6.0977
3.3085 | -34.1397 | -6.1093
3.3146 | -34.1398 | -6.1205
3.3207 | -34.1401 | -6.132
3.3268 | -34.1403 | -6.1435
3.3329 | -34.1405 | -6.1542
3.339 | -34.1407 | -6.1651
3.3451 | -34.1409 | -6.1763
3.3512 | -34.1411 | -6.187
3.3573 | -34.1413 | -6.1983
3.3634 | -34.1415 | -6.2094
3.3696 | -34.1417 | -6.2208
3.3757 | -34.1418 | -6.2316
3.3818 | -34.1421 | -6.2426
3.3879 | -34.1423 | -6.2537
3.394 | -34.1424 | -6.2651
3.4001 | -34.1426 | -6.276
3.4062 | -34.1427 | -6.287
3.4123 | -34.1429 | -6.2976
3.4184 | -34.1431 | -6.309
3.4245 | -34.1433 | -6.3198
3.4306 | -34.1434 | -6.3305
3.4367 | -34.1436 | -6.3412
3.4428 | -34.1438 | -6.3517
3.4489 | -34.144 | -6.3626
3.455 | -34.1441 | -6.3736
3.4611 | -34.1443 | -6.3849
3.4672 | -34.1445 | -6.396
3.4733 | -34.1447 | -6.4071
3.4794 | -34.1449 | -6.418
3.4855 | -34.145 | -6.4286
3.4916 | -34.1452 | -6.439
3.4977 | -34.1454 | -6.4504
3.5038 | -34.1455 | -6.4613
3.5099 | -34.1457 | -6.4717
3.5161 | -34.1459 | -6.4822
3.5222 | -34.146 | -6.4927
3.5283 | -34.1461 | -6.5036
3.5344 | -34.1462 | -6.5145
3.5405 | -34.1464 | -6.5252
3.5466 | -34.1466 | -6.5362
3.5527 | -34.1467 | -6.5472
3.5588 | -34.1469 | -6.5579
3.5649 | -34.1469 | -6.5687
3.571 | -34.1471 | -6.5795
3.5771 | -34.1472 | -6.5906
3.5832 | -34.1474 | -6.6007
3.5893 | -34.1476 | -6.6114
3.5954 | -34.1478 | -6.6224
3.6015 | -34.148 | -6.6334
3.6076 | -34.1482 | -6.644
3.6137 | -34.1484 | -6.6543
3.6198 | -34.1486 | -6.6647
3.6259 | -34.1488 | -6.6757
3.632 | -34.149 | -6.6865
3.6381 | -34.1492 | -6.6974
3.6442 | -34.1493 | -6.7082
3.6503 | -34.1495 | -6.7185
3.6565 | -34.1498 | -6.7294
3.6626 | -34.1499 | -6.7401
3.6687 | -34.1502 | -6.7509
3.6748 | -34.1504 | -6.7614
3.6809 | -34.1507 | -6.7722
3.687 | -34.1508 | -6.7833
3.6931 | -34.1511 | -6.7937
3.6992 | -34.1513 | -6.8045
3.7053 | -34.1515 | -6.8152
3.7114 | -34.1517 | -6.8261
3.7175 | -34.1519 | -6.8369
3.7236 | -34.1522 | -6.8472
3.7297 | -34.1524 | -6.8581
3.7358 | -34.1526 | -6.8691
3.7419 | -34.1528 | -6.8801
3.748 | -34.1531 | -6.8907
3.7541 | -34.1533 | -6.9015
3.7602 | -34.1536 | -6.9122
3.7663 | -34.1538 | -6.9228
3.7724 | -34.154 | -6.9337
3.7785 | -34.1543 | -6.9443
3.7846 | -34.1545 | -6.9549
3.7907 | -34.1548 | -6.966
3.7969 | -34.1551 | -6.9768
3.803 | -34.1553 | -6.9872
3.8091 | -34.1555 | -6.9979
3.8152 | -34.1558 | -7.0087
3.8213 | -34.1561 | -7.0197
3.8274 | -34.1563 | -7.0301
3.8335 | -34.1567 | -7.0411
3.8396 | -34.157 | -7.0519
3.8457 | -34.1573 | -7.0628
3.8518 | -34.1576 | -7.0734
3.8579 | -34.1579 | -7.0842
3.864 | -34.1582 | -7.0951
3.8701 | -34.1585 | -7.1062
3.8762 | -34.1587 | -7.1168
3.8823 | -34.159 | -7.1278
3.8884 | -34.1593 | -7.1386
3.8945 | -34.1597 | -7.1496
3.9006 | -34.1599 | -7.1607
3.9067 | -34.1602 | -7.172
3.9128 | -34.1605 | -7.1829
3.9189 | -34.1608 | -7.1936
3.925 | -34.161 | -7.2045
3.9311 | -34.1613 | -7.2149
3.9372 | -34.1616 | -7.2259
3.9434 | -34.1619 | -7.2367
3.9495 | -34.1622 | -7.2477
3.9556 | -34.1626 | -7.2583
3.9617 | -34.1628 | -7.2687
3.9678 | -34.1632 | -7.2793
3.9739 | -34.1634 | -7.2901
3.98 | -34.1637 | -7.3009
3.9861 | -34.164 | -7.3118
3.9922 | -34.1642 | -7.3226
3.9983 | -34.1645 | -7.333
4.0044 | -34.1648 | -7.3434
4.0105 | -34.165 | -7.3542
4.0166 | -34.1654 | -7.3651
4.0227 | -34.1656 | -7.3758
4.0288 | -34.1659 | -7.387
4.0349 | -34.1661 | -7.3975
4.041 | -34.1664 | -7.4083
4.0471 | -34.1666 | -7.4191
4.0532 | -34.1669 | -7.4296
4.0593 | -34.1672 | -7.4405
4.0654 | -34.1675 | -7.4511
4.0715 | -34.1677 | -7.4623
4.0776 | -34.168 | -7.473
4.0838 | -34.1683 | -7.4841
4.0899 | -34.1686 | -7.4947
4.096 | -34.1689 | -7.5054
4.1021 | -34.1691 | -7.5159
4.1082 | -34.1694 | -7.5264
4.1143 | -34.1696 | -7.5368
4.1204 | -34.1699 | -7.5473
4.1265 | -34.1702 | -7.5577
4.1326 | -34.1704 | -7.5682
4.1387 | -34.1707 | -7.5788
4.1448 | -34.171 | -7.5895
4.1509 | -34.1712 | -7.6002
4.157 | -34.1714 | -7.6113
4.1631 | -34.1717 | -7.6222
4.1692 | -34.172 | -7.6327
4.1753 | -34.1722 | -7.6434
4.1814 | -34.1724 | -7.6537
4.1875 | -34.1727 | -7.6646
4.1936 | -34.173 | -7.6754
4.1997 | -34.1732 | -7.6857
4.2058 | -34.1734 | -7.6966
4.2119 | -34.1736 | -7.7071
4.218 | -34.1739 | -7.7175
4.2241 | -34.1741 | -7.7279
4.2303 | -34.1744 | -7.7384
4.2364 | -34.1746 | -7.749
4.2425 | -34.1748 | -7.7596
4.2486 | -34.175 | -7.7704
4.2547 | -34.1751 | -7.7808
4.2608 | -34.1753 | -7.7915
4.2669 | -34.1755 | -7.8018
4.273 | -34.1758 | -7.8125
4.2791 | -34.176 | -7.823
4.2852 | -34.1762 | -7.8338
4.2913 | -34.1764 | -7.8441
4.2974 | -34.1766 | -7.8549
4.3035 | -34.1768 | -7.8658
4.3096 | -34.177 | -7.8765
4.3157 | -34.1772 | -7.8869
4.3218 | -34.1774 | -7.8973
4.3279 | -34.1777 | -7.9079
4.334 | -34.1778 | -7.9184
4.3401 | -34.178 | -7.9285
4.3462 | -34.1783 | -7.9392
4.3523 | -34.1786 | -7.9493
4.3584 | -34.1787 | -7.9595
4.3645 | -34.1789 | -7.9699
4.3707 | -34.1791 | -7.9802
4.3768 | -34.1793 | -7.9906
4.3829 | -34.1796 | -8.0014
4.389 | -34.1798 | -8.0118
4.3951 | -34.18 | -8.0221
4.4012 | -34.1801 | -8.0329
4.4073 | -34.1803 | -8.0432
4.4134 | -34.1805 | -8.0536
4.4195 | -34.1808 | -8.064
4.4256 | -34.181 | -8.0743
4.4317 | -34.1812 | -8.0846
4.4378 | -34.1814 | -8.0947
4.4439 | -34.1816 | -8.1048
4.45 | -34.1818 | -8.1152
4.4561 | -34.1819 | -8.1256
4.4622 | -34.1821 | -8.1361
4.4683 | -34.1823 | -8.1464
4.4744 | -34.1825 | -8.1571
4.4805 | -34.1827 | -8.1672
4.4866 | -34.1829 | -8.1775
4.4927 | -34.1832 | -8.1881
4.4988 | -34.1834 | -8.1985
4.5049 | -34.1836 | -8.2085
4.511 | -34.1839 | -8.2186
4.5172 | -34.1841 | -8.229
4.5233 | -34.1842 | -8.2397
4.5294 | -34.1845 | -8.2499
4.5355 | -34.1847 | -8.2603
4.5416 | -34.1849 | -8.2706
4.5477 | -34.1851 | -8.2809
4.5538 | -34.1853 | -8.291
4.5599 | -34.1856 | -8.3012
4.566 | -34.1858 | -8.3116
4.5721 | -34.186 | -8.322
4.5782 | -34.1863 | -8.3326
4.5843 | -34.1865 | -8.343
4.5904 | -34.1867 | -8.3534
4.5965 | -34.187 | -8.3636
4.6026 | -34.1872 | -8.3743
4.6087 | -34.1874 | -8.3846
4.6148 | -34.1876 | -8.3953
4.6209 | -34.1878 | -8.4056
4.627 | -34.1881 | -8.4161
4.6331 | -34.1882 | -8.4263
4.6392 | -34.1885 | -8.4367
4.6453 | -34.1887 | -8.4471
4.6514 | -34.189 | -8.4575
4.6576 | -34.1892 | -8.4682
4.6637 | -34.1893 | -8.4783
4.6698 | -34.1895 | -8.4885
4.6759 | -34.1898 | -8.4992
4.682 | -34.19 | -8.5095
4.6881 | -34.1901 | -8.5198
4.6942 | -34.1904 | -8.5302
4.7003 | -34.1906 | -8.5406
4.7064 | -34.1908 | -8.5508
4.7125 | -34.1911 | -8.5615
4.7186 | -34.1913 | -8.5717
4.7247 | -34.1916 | -8.5821
4.7308 | -34.1918 | -8.5925
4.7369 | -34.192 | -8.6029
4.743 | -34.1923 | -8.6137
4.7491 | -34.1925 | -8.6244
4.7552 | -34.1927 | -8.6349
4.7613 | -34.1929 | -8.6454
4.7674 | -34.1931 | -8.6557
4.7735 | -34.1933 | -8.6665
4.7796 | -34.1937 | -8.6772
4.7857 | -34.1939 | -8.6877
4.7918 | -34.1941 | -8.6979
4.7979 | -34.1944 | -8.7082
4.8041 | -34.1946 | -8.7187
4.8102 | -34.1949 | -8.7288
4.8163 | -34.1951 | -8.7392
4.8224 | -34.1954 | -8.7496
4.8285 | -34.1956 | -8.7598
4.8346 | -34.1959 | -8.77
4.8407 | -34.1961 | -8.7806
4.8468 | -34.1964 | -8.7913
4.8529 | -34.1966 | -8.802
4.859 | -34.1969 | -8.8127
4.8651 | -34.1971 | -8.823
4.8712 | -34.1974 | -8.8333
4.8773 | -34.1976 | -8.8438
4.8834 | -34.1978 | -8.8542
4.8895 | -34.1981 | -8.8645
4.8956 | -34.1983 | -8.875
4.9017 | -34.1985 | -8.8853
4.9078 | -34.1988 | -8.896
4.9139 | -34.199 | -8.9061
4.92 | -34.1993 | -8.9163
4.9261 | -34.1996 | -8.9264
4.9322 | -34.1999 | -8.9368
4.9383 | -34.2002 | -8.9472
4.9445 | -34.2004 | -8.9577
4.9506 | -34.2006 | -8.968
4.9567 | -34.2009 | -8.9782
4.9628 | -34.2011 | -8.9887
4.9689 | -34.2014 | -8.9995
4.975 | -34.2016 | -9.0102
4.9811 | -34.2019 | -9.0203
4.9872 | -34.2022 | -9.0307
4.9933 | -34.2026 | -9.0414
4.9994 | -34.2028 | -9.0522
5.0055 | -34.2031 | -9.0627
5.0116 | -34.2034 | -9.0735
5.0177 | -34.2037 | -9.0842
5.0238 | -34.204 | -9.095
5.0299 | -34.2044 | -9.1058
5.036 | -34.2047 | -9.1162
5.0421 | -34.2049 | -9.1266
5.0482 | -34.2052 | -9.1371
5.0543 | -34.2055 | -9.1473
5.0604 | -34.2058 | -9.1578
5.0665 | -34.2061 | -9.168
5.0726 | -34.2064 | -9.1788
5.0787 | -34.2067 | -9.1895
5.0848 | -34.207 | -9.2001
5.091 | -34.2072 | -9.2104
5.0971 | -34.2076 | -9.2206
5.1032 | -34.2079 | -9.2314
5.1093 | -34.2082 | -9.2419
5.1154 | -34.2084 | -9.2522
5.1215 | -34.2087 | -9.2625
5.1276 | -34.209 | -9.2728
5.1337 | -34.2093 | -9.283
5.1398 | -34.2096 | -9.2933
5.1459 | -34.2099 | -9.304
5.152 | -34.2102 | -9.3146
5.1581 | -34.2105 | -9.3253
5.1642 | -34.2108 | -9.3358
5.1703 | -34.2111 | -9.3462
5.1764 | -34.2114 | -9.3565
5.1825 | -34.2117 | -9.3668
5.1886 | -34.212 | -9.3769
5.1947 | -34.2123 | -9.3874
5.2008 | -34.2126 | -9.3979
5.2069 | -34.2129 | -9.4086
5.213 | -34.2132 | -9.4188
5.2191 | -34.2135 | -9.4294
5.2252 | -34.2138 | -9.4399
5.2314 | -34.2141 | -9.4506
5.2375 | -34.2144 | -9.4608
5.2436 | -34.2147 | -9.4713
5.2497 | -34.215 | -9.4815
5.2558 | -34.2153 | -9.4923
5.2619 | -34.2156 | -9.5031
5.268 | -34.2159 | -9.5137
5.2741 | -34.2163 | -9.5242
5.2802 | -34.2166 | -9.5351
5.2863 | -34.2169 | -9.5455
5.2924 | -34.2172 | -9.5561
5.2985 | -34.2176 | -9.5667
5.3046 | -34.2179 | -9.5773
5.3107 | -34.2182 | -9.5879
5.3168 | -34.2185 | -9.5985
5.3229 | -34.2188 | -9.609
5.329 | -34.2191 | -9.6195
5.3351 | -34.2194 | -9.6298
5.3412 | -34.2197 | -9.6402
5.3473 | -34.22 | -9.6508
5.3534 | -34.2204 | -9.6614
5.3595 | -34.2207 | -9.6719
5.3656 | -34.221 | -9.6824
5.3717 | -34.2213 | -9.6931
5.3779 | -34.2216 | -9.704
5.384 | -34.2219 | -9.7145
5.3901 | -34.2223 | -9.7253
5.3962 | -34.2225 | -9.7356
5.4023 | -34.2229 | -9.7463
5.4084 | -34.2233 | -9.7567
5.4145 | -34.2236 | -9.767
5.4206 | -34.2239 | -9.7776
5.4267 | -34.2242 | -9.7884
5.4328 | -34.2245 | -9.799
5.4389 | -34.2248 | -9.8092
5.445 | -34.2251 | -9.82
5.4511 | -34.2254 | -9.8304
5.4572 | -34.2258 | -9.841
5.4633 | -34.2261 | -9.8516
5.4694 | -34.2264 | -9.8621
5.4755 | -34.2267 | -9.8724
5.4816 | -34.227 | -9.883
5.4877 | -34.2273 | -9.8933
5.4938 | -34.2276 | -9.9036
5.4999 | -34.228 | -9.914
5.506 | -34.2283 | -9.9245
5.5121 | -34.2286 | -9.9349
5.5183 | -34.2289 | -9.9451
5.5244 | -34.2292 | -9.9558
5.5305 | -34.2295 | -9.9662
5.5366 | -34.2298 | -9.9765
5.5427 | -34.2301 | -9.9871
5.5488 | -34.2304 | -9.9977
5.5549 | -34.2307 | -10.0081
5.561 | -34.2311 | -10.0189
5.5671 | -34.2314 | -10.0292
5.5732 | -34.2317 | -10.0399
5.5793 | -34.232 | -10.0507
5.5854 | -34.2324 | -10.0611
5.5915 | -34.2327 | -10.0719
5.5976 | -34.233 | -10.0825
5.6037 | -34.2333 | -10.093
5.6098 | -34.2336 | -10.1035
5.6159 | -34.2339 | -10.1138
5.622 | -34.2342 | -10.1244
5.6281 | -34.2345 | -10.1349
5.6342 | -34.2348 | -10.1457
5.6403 | -34.2351 | -10.1563
5.6464 | -34.2354 | -10.1669
5.6525 | -34.2357 | -10.1775
5.6586 | -34.2361 | -10.1877
5.6648 | -34.2364 | -10.1983
5.6709 | -34.2368 | -10.2089
5.677 | -34.2371 | -10.2195
5.6831 | -34.2374 | -10.23
5.6892 | -34.2377 | -10.2407
5.6953 | -34.238 | -10.2512
5.7014 | -34.2383 | -10.2618
5.7075 | -34.2386 | -10.2723
5.7136 | -34.2389 | -10.2827
5.7197 | -34.2392 | -10.2932
5.7258 | -34.2394 | -10.3035
5.7319 | -34.2398 | -10.314
5.738 | -34.2401 | -10.3245
5.7441 | -34.2404 | -10.335
5.7502 | -34.2408 | -10.3452
5.7563 | -34.2411 | -10.3557
5.7624 | -34.2414 | -10.3664
5.7685 | -34.2417 | -10.3769
5.7746 | -34.242 | -10.3877
5.7807 | -34.2423 | -10.3982
5.7868 | -34.2425 | -10.4089
5.7929 | -34.2428 | -10.4193
5.799 | -34.2431 | -10.4296
5.8052 | -34.2434 | -10.44
5.8113 | -34.2437 | -10.4504
5.8174 | -34.244 | -10.4607
5.8235 | -34.2444 | -10.4712
5.8296 | -34.2446 | -10.4818
5.8357 | -34.2449 | -10.4922
5.8418 | -34.2452 | -10.5028
5.8479 | -34.2455 | -10.5133
5.854 | -34.2458 | -10.524
5.8601 | -34.2461 | -10.5343
5.8662 | -34.2464 | -10.545
5.8723 | -34.2467 | -10.5553
5.8784 | -34.247 | -10.5657
5.8845 | -34.2473 | -10.5761
5.8906 | -34.2476 | -10.5866
5.8967 | -34.2479 | -10.5969
5.9028 | -34.2482 | -10.6072
5.9089 | -34.2486 | -10.6175
5.915 | -34.2489 | -10.6278
5.9211 | -34.2492 | -10.6382
5.9272 | -34.2495 | -10.6484
5.9333 | -34.2498 | -10.6587
5.9394 | -34.2501 | -10.6691
5.9455 | -34.2504 | -10.6793
5.9517 | -34.2508 | -10.6896
5.9578 | -34.2511 | -10.7001
5.9639 | -34.2514 | -10.7102
5.97 | -34.2518 | -10.7208
5.9761 | -34.2521 | -10.731
5.9822 | -34.2524 | -10.7416
5.9883 | -34.2527 | -10.752
5.9944 | -34.253 | -10.7625
6.0005 | -34.2534 | -10.7731
6.0066 | -34.2537 | -10.7833
6.0127 | -34.254 | -10.7934
6.0188 | -34.2543 | -10.8041
6.0249 | -34.2546 | -10.8145
6.031 | -34.2549 | -10.8247
6.0371 | -34.2552 | -10.8351
6.0432 | -34.2555 | -10.8454
6.0493 | -34.2559 | -10.8557
6.0554 | -34.2562 | -10.8659
6.0615 | -34.2565 | -10.8763
6.0676 | -34.2569 | -10.8869
6.0737 | -34.2572 | -10.8974
6.0798 | -34.2575 | -10.9078
6.0859 | -34.2578 | -10.9181
6.0921 | -34.2581 | -10.9283
6.0982 | -34.2584 | -10.9386
6.1043 | -34.2588 | -10.9489
6.1104 | -34.2591 | -10.9595
6.1165 | -34.2595 | -10.9694
6.1226 | -34.2598 | -10.9797
6.1287 | -34.2601 | -10.99
6.1348 | -34.2604 | -11.0004
6.1409 | -34.2607 | -11.0107
6.147 | -34.261 | -11.021
6.1531 | -34.2614 | -11.0313
6.1592 | -34.2617 | -11.0415
6.1653 | -34.262 | -11.0517
6.1714 | -34.2623 | -11.0621
6.1775 | -34.2626 | -11.0726
6.1836 | -34.2629 | -11.0829
6.1897 | -34.2633 | -11.0933
6.1958 | -34.2636 | -11.1039
6.2019 | -34.2639 | -11.1142
6.208 | -34.2643 | -11.1246
6.2141 | -34.2646 | -11.1349
6.2202 | -34.2649 | -11.1453
6.2263 | -34.2653 | -11.1556
6.2325 | -34.2656 | -11.1657
6.2386 | -34.2659 | -11.1757
6.2447 | -34.2662 | -11.1863
6.2508 | -34.2666 | -11.1966
6.2569 | -34.2669 | -11.2069
6.263 | -34.2672 | -11.2172
6.2691 | -34.2675 | -11.2275
6.2752 | -34.2678 | -11.2379
6.2813 | -34.2681 | -11.2481
6.2874 | -34.2685 | -11.2583
6.2935 | -34.2688 | -11.2686
6.2996 | -34.2691 | -11.2788
6.3057 | -34.2694 | -11.2892
6.3118 | -34.2697 | -11.2996
6.3179 | -34.27 | -11.3099
6.324 | -34.2703 | -11.3203
6.3301 | -34.2707 | -11.3306
6.3362 | -34.271 | -11.3408
6.3423 | -34.2713 | -11.3513
6.3484 | -34.2717 | -11.3617
6.3545 | -34.272 | -11.3717
6.3606 | -34.2723 | -11.3818
6.3667 | -34.2726 | -11.392
6.3728 | -34.2729 | -11.4023
6.379 | -34.2732 | -11.4126
6.3851 | -34.2736 | -11.4226
6.3912 | -34.2738 | -11.4329
6.3973 | -34.2742 | -11.4432
6.4034 | -34.2745 | -11.4533
6.4095 | -34.2748 | -11.4636
6.4156 | -34.2751 | -11.4739
6.4217 | -34.2754 | -11.4842
6.4278 | -34.2757 | -11.4944
6.4339 | -34.2761 | -11.5047
6.44 | -34.2764 | -11.5148
6.4461 | -34.2767 | -11.5253
6.4522 | -34.277 | -11.5354
6.4583 | -34.2773 | -11.5457
6.4644 | -34.2776 | -11.5561
6.4705 | -34.278 | -11.5661
6.4766 | -34.2783 | -11.5764
6.4827 | -34.2786 | -11.5867
6.4888 | -34.2789 | -11.5968
6.4949 | -34.2792 | -11.6068
6.501 | -34.2795 | -11.6171
6.5071 | -34.2799 | -11.6272
6.5132 | -34.2802 | -11.6375
6.5194 | -34.2805 | -11.6478
6.5255 | -34.2808 | -11.6579
6.5316 | -34.2811 | -11.6682
6.5377 | -34.2814 | -11.6785
6.5438 | -34.2817 | -11.6886
6.5499 | -34.2821 | -11.6987
6.556 | -34.2824 | -11.7089
6.5621 | -34.2827 | -11.7194
6.5682 | -34.283 | -11.7297
6.5743 | -34.2834 | -11.7399
6.5804 | -34.2837 | -11.7502
6.5865 | -34.284 | -11.7605
6.5926 | -34.2843 | -11.7706
6.5987 | -34.2846 | -11.7806
6.6048 | -34.285 | -11.7909
6.6109 | -34.2853 | -11.8012
6.617 | -34.2856 | -11.8116
6.6231 | -34.2859 | -11.8218
6.6292 | -34.2862 | -11.8321
6.6353 | -34.2866 | -11.8426
6.6414 | -34.2869 | -11.8528
6.6475 | -34.2872 | -11.863
6.6536 | -34.2875 | -11.8732
6.6597 | -34.2878 | -11.8832
6.6659 | -34.2882 | -11.8935
6.672 | -34.2885 | -11.9036
6.6781 | -34.2888 | -11.9139
6.6842 | -34.2892 | -11.924
6.6903 | -34.2895 | -11.9341
6.6964 | -34.2898 | -11.9443
6.7025 | -34.2901 | -11.9543
6.7086 | -34.2905 | -11.9646
6.7147 | -34.2908 | -11.9748
6.7208 | -34.2911 | -11.9852
6.7269 | -34.2914 | -11.9953
6.733 | -34.2917 | -12.0053
6.7391 | -34.292 | -12.0156
6.7452 | -34.2924 | -12.0255
6.7513 | -34.2927 | -12.0355
6.7574 | -34.293 | -12.0458
6.7635 | -34.2933 | -12.0558
6.7696 | -34.2937 | -12.0659
6.7757 | -34.294 | -12.0761
6.7818 | -34.2943 | -12.0864
6.7879 | -34.2947 | -12.0964
6.794 | -34.295 | -12.1064
6.8001 | -34.2953 | -12.1166
6.8063 | -34.2956 | -12.1268
6.8124 | -34.2959 | -12.1371
6.8185 | -34.2962 | -12.1475
6.8246 | -34.2965 | -12.1577
6.8307 | -34.2969 | -12.1677
6.8368 | -34.2972 | -12.1781
6.8429 | -34.2975 | -12.1884
6.849 | -34.2979 | -12.1986
6.8551 | -34.2982 | -12.2086
6.8612 | -34.2985 | -12.2188
6.8673 | -34.2988 | -12.2289
6.8734 | -34.2992 | -12.2393
6.8795 | -34.2995 | -12.2494
6.8856 | -34.2998 | -12.2596
6.8917 | -34.3001 | -12.2697
6.8978 | -34.3005 | -12.2797
6.9039 | -34.3008 | -12.2896
6.91 | -34.3011 | -12.2998
6.9161 | -34.3015 | -12.3098
6.9222 | -34.3018 | -12.3199
6.9283 | -34.3021 | -12.3298
6.9344 | -34.3025 | -12.34
6.9405 | -34.3028 | -12.35
6.9466 | -34.3032 | -12.3604
6.9528 | -34.3035 | -12.3707
6.9589 | -34.3039 | -12.381
6.965 | -34.3042 | -12.3912
6.9711 | -34.3046 | -12.4015
6.9772 | -34.3049 | -12.4116
6.9833 | -34.3052 | -12.4217
6.9894 | -34.3056 | -12.4321
6.9955 | -34.3059 | -12.442
7.0016 | -34.3063 | -12.4521
7.0077 | -34.3066 | -12.4622
7.0138 | -34.3069 | -12.4723
7.0199 | -34.3073 | -12.4824
7.026 | -34.3076 | -12.4925
7.0321 | -34.3079 | -12.5024
7.0382 | -34.3083 | -12.5123
7.0443 | -34.3086 | -12.5226
7.0504 | -34.309 | -12.5327
7.0565 | -34.3093 | -12.5428
7.0626 | -34.3096 | -12.5527
7.0687 | -34.31 | -12.5625
7.0748 | -34.3103 | -12.5723
7.0809 | -34.3107 | -12.5824
7.087 | -34.311 | -12.5926
7.0932 | -34.3113 | -12.6028
7.0993 | -34.3116 | -12.6128
7.1054 | -34.312 | -12.6228
7.1115 | -34.3124 | -12.6327
7.1176 | -34.3127 | -12.6427
7.1237 | -34.313 | -12.6527
7.1298 | -34.3133 | -12.6627
7.1359 | -34.3137 | -12.6728
7.142 | -34.314 | -12.6826
7.1481 | -34.3143 | -12.6925
7.1542 | -34.3146 | -12.7026
7.1603 | -34.315 | -12.7127
7.1664 | -34.3154 | -12.7227
7.1725 | -34.3157 | -12.7327
7.1786 | -34.3161 | -12.7427
7.1847 | -34.3164 | -12.7529
7.1908 | -34.3167 | -12.7631
7.1969 | -34.317 | -12.7732
7.203 | -34.3174 | -12.7831
7.2091 | -34.3177 | -12.793
7.2152 | -34.3181 | -12.803
7.2213 | -34.3184 | -12.8128
7.2274 | -34.3188 | -12.8229
7.2335 | -34.3191 | -12.8328
7.2397 | -34.3195 | -12.8428
7.2458 | -34.3198 | -12.8525
7.2519 | -34.3202 | -12.8625
7.258 | -34.3205 | -12.8727
7.2641 | -34.3209 | -12.8827
7.2702 | -34.3212 | -12.8929
7.2763 | -34.3215 | -12.903
7.2824 | -34.3219 | -12.9131
7.2885 | -34.3223 | -12.9231
7.2946 | -34.3226 | -12.9331
7.3007 | -34.3229 | -12.9432
7.3068 | -34.3232 | -12.9534
7.3129 | -34.3236 | -12.9634
7.319 | -34.324 | -12.9734
7.3251 | -34.3243 | -12.9832
7.3312 | -34.3246 | -12.9933
7.3373 | -34.325 | -13.0031
7.3434 | -34.3254 | -13.013
7.3495 | -34.3257 | -13.0228
7.3556 | -34.3261 | -13.0329
7.3617 | -34.3264 | -13.043
7.3678 | -34.3268 | -13.0529
7.3739 | -34.3271 | -13.0631
7.3801 | -34.3274 | -13.0733
7.3862 | -34.3278 | -13.0832
7.3923 | -34.3282 | -13.0933
7.3984 | -34.3285 | -13.1031
7.4045 | -34.3289 | -13.113
7.4106 | -34.3293 | -13.1229
7.4167 | -34.3296 | -13.133
7.4228 | -34.33 | -13.1431
7.4289 | -34.3303 | -13.153
7.435 | -34.3307 | -13.1628
7.4411 | -34.331 | -13.1728
7.4472 | -34.3313 | -13.1828
7.4533 | -34.3317 | -13.1925
7.4594 | -34.3321 | -13.2027
7.4655 | -34.3324 | -13.2126
7.4716 | -34.3328 | -13.2227
7.4777 | -34.3331 | -13.2327
7.4838 | -34.3335 | -13.2424
7.4899 | -34.3338 | -13.2522
7.496 | -34.3341 | -13.262
7.5021 | -34.3345 | -13.2719
7.5082 | -34.3349 | -13.2819
7.5143 | -34.3352 | -13.2919
7.5204 | -34.3355 | -13.302
7.5266 | -34.3359 | -13.3117
7.5327 | -34.3362 | -13.3216
7.5388 | -34.3366 | -13.3318
7.5449 | -34.337 | -13.3419
7.551 | -34.3373 | -13.3517
7.5571 | -34.3377 | -13.3618
7.5632 | -34.338 | -13.3717
7.5693 | -34.3384 | -13.3816
7.5754 | -34.3388 | -13.3913
7.5815 | -34.3391 | -13.4013
7.5876 | -34.3395 | -13.4115
7.5937 | -34.3398 | -13.4215
7.5998 | -34.3402 | -13.4314
7.6059 | -34.3406 | -13.4415
7.612 | -34.3409 | -13.4515
7.6181 | -34.3413 | -13.4613
7.6242 | -34.3417 | -13.4715
7.6303 | -34.342 | -13.4816
7.6364 | -34.3424 | -13.4916
7.6425 | -34.3427 | -13.5016
7.6486 | -34.343 | -13.5116
7.6547 | -34.3434 | -13.5214
7.6608 | -34.3438 | -13.5313
7.667 | -34.3441 | -13.5413
7.6731 | -34.3445 | -13.5513
7.6792 | -34.3448 | -13.5615
7.6853 | -34.3452 | -13.5716
7.6914 | -34.3456 | -13.5816
7.6975 | -34.3459 | -13.5916
7.7036 | -34.3463 | -13.6016
7.7097 | -34.3467 | -13.6117
7.7158 | -34.347 | -13.6217
7.7219 | -34.3474 | -13.6315
7.728 | -34.3477 | -13.6415
7.7341 | -34.3481 | -13.6515
7.7402 | -34.3485 | -13.6615
7.7463 | -34.3488 | -13.6716
7.7524 | -34.3492 | -13.6818
7.7585 | -34.3495 | -13.6917
7.7646 | -34.3499 | -13.7017
7.7707 | -34.3502 | -13.7118
7.7768 | -34.3506 | -13.7219
7.7829 | -34.3509 | -13.7321
7.789 | -34.3513 | -13.7423
7.7951 | -34.3516 | -13.7523
7.8012 | -34.352 | -13.7623
7.8073 | -34.3524 | -13.7724
7.8135 | -34.3528 | -13.7824
7.8196 | -34.3532 | -13.7925
7.8257 | -34.3535 | -13.8027
7.8318 | -34.3538 | -13.8128
7.8379 | -34.3542 | -13.8227
7.844 | -34.3546 | -13.8326
7.8501 | -34.3549 | -13.8428
7.8562 | -34.3553 | -13.853
7.8623 | -34.3557 | -13.8629
7.8684 | -34.356 | -13.873
7.8745 | -34.3564 | -13.8831
7.8806 | -34.3567 | -13.8932
7.8867 | -34.3571 | -13.903
7.8928 | -34.3575 | -13.9128
7.8989 | -34.3578 | -13.9228
7.905 | -34.3582 | -13.9329
7.9111 | -34.3585 | -13.943
7.9172 | -34.3589 | -13.953
7.9233 | -34.3592 | -13.9631
7.9294 | -34.3596 | -13.9732
7.9355 | -34.3599 | -13.983
7.9416 | -34.3603 | -13.9931
7.9477 | -34.3606 | -14.0029
7.9539 | -34.361 | -14.0128
7.96 | -34.3613 | -14.0229
7.9661 | -34.3617 | -14.0329
7.9722 | -34.3621 | -14.0428
7.9783 | -34.3624 | -14.053
7.9844 | -34.3628 | -14.0629
7.9905 | -34.3632 | -14.0731
7.9966 | -34.3635 | -14.0832
8.0027 | -34.3639 | -14.0932
8.0088 | -34.3642 | -14.1032
8.0149 | -34.3646 | -14.1133
8.021 | -34.3649 | -14.1233
8.0271 | -34.3653 | -14.1336
8.0332 | -34.3656 | -14.1435
8.0393 | -34.366 | -14.1536
8.0454 | -34.3664 | -14.1638
8.0515 | -34.3668 | -14.1738
8.0576 | -34.3672 | -14.1841
8.0637 | -34.3675 | -14.194
8.0698 | -34.3678 | -14.2039
8.0759 | -34.3682 | -14.2138
8.082 | -34.3686 | -14.2238
8.0881 | -34.3689 | -14.2338
8.0942 | -34.3693 | -14.2439
8.1004 | -34.3697 | -14.254
8.1065 | -34.37 | -14.2639
8.1126 | -34.3704 | -14.274
8.1187 | -34.3707 | -14.2841
8.1248 | -34.3711 | -14.2939
8.1309 | -34.3715 | -14.3038
8.137 | -34.3718 | -14.3136
8.1431 | -34.3722 | -14.3238
8.1492 | -34.3725 | -14.3338
8.1553 | -34.3729 | -14.3439
8.1614 | -34.3733 | -14.3541
8.1675 | -34.3736 | -14.3641
8.1736 | -34.3739 | -14.3743
8.1797 | -34.3743 | -14.3844
8.1858 | -34.3747 | -14.3944
8.1919 | -34.375 | -14.4044
8.198 | -34.3754 | -14.4143
8.2041 | -34.3758 | -14.4244
8.2102 | -34.3761 | -14.4344
8.2163 | -34.3765 | -14.4443
8.2224 | -34.3768 | -14.4543
8.2285 | -34.3772 | -14.4641
8.2346 | -34.3776 | -14.474
8.2408 | -34.3779 | -14.4839
8.2469 | -34.3783 | -14.494
8.253 | -34.3787 | -14.5039
8.2591 | -34.379 | -14.514
8.2652 | -34.3794 | -14.5241
8.2713 | -34.3798 | -14.5342
8.2774 | -34.3801 | -14.5443
8.2835 | -34.3805 | -14.5545
8.2896 | -34.3809 | -14.5643
8.2957 | -34.3812 | -14.5744
8.3018 | -34.3816 | -14.5845
8.3079 | -34.382 | -14.5944
8.314 | -34.3823 | -14.6044
8.3201 | -34.3827 | -14.6143
8.3262 | -34.3831 | -14.6244
8.3323 | -34.3834 | -14.6343
8.3384 | -34.3838 | -14.6444
8.3445 | -34.3841 | -14.6543
8.3506 | -34.3845 | -14.6644
8.3567 | -34.3849 | -14.6744
8.3628 | -34.3852 | -14.6844
8.3689 | -34.3856 | -14.6943
8.375 | -34.386 | -14.7043
8.3812 | -34.3863 | -14.7142
8.3873 | -34.3867 | -14.7242
8.3934 | -34.3871 | -14.7343
8.3995 | -34.3874 | -14.7441
8.4056 | -34.3878 | -14.7541
8.4117 | -34.3881 | -14.7642
8.4178 | -34.3885 | -14.7742
8.4239 | -34.3888 | -14.784
8.43 | -34.3892 | -14.7941
8.4361 | -34.3895 | -14.8042
8.4422 | -34.3899 | -14.814
8.4483 | -34.3903 | -14.8239
8.4544 | -34.3906 | -14.8339
8.4605 | -34.391 | -14.8438
8.4666 | -34.3914 | -14.8539
8.4727 | -34.3917 | -14.8639
8.4788 | -34.3921 | -14.8739
8.4849 | -34.3924 | -14.8838
8.491 | -34.3928 | -14.8938
8.4971 | -34.3932 | -14.9036
8.5032 | -34.3935 | -14.9135
8.5093 | -34.3939 | -14.9233
8.5154 | -34.3943 | -14.9333
8.5215 | -34.3947 | -14.9431
8.5277 | -34.395 | -14.9531
8.5338 | -34.3954 | -14.963
8.5399 | -34.3957 | -14.9729
8.546 | -34.3961 | -14.9828
8.5521 | -34.3965 | -14.9927
8.5582 | -34.3968 | -15.0027
8.5643 | -34.3972 | -15.0124
8.5704 | -34.3975 | -15.0221
8.5765 | -34.3979 | -15.0322
8.5826 | -34.3983 | -15.0418
8.5887 | -34.3987 | -15.0517
8.5948 | -34.399 | -15.0616
8.6009 | -34.3994 | -15.0713
8.607 | -34.3998 | -15.0814
8.6131 | -34.4001 | -15.0913
8.6192 | -34.4005 | -15.1012
8.6253 | -34.4009 | -15.111
8.6314 | -34.4013 | -15.1209
8.6375 | -34.4016 | -15.1308
8.6436 | -34.402 | -15.1407
8.6497 | -34.4023 | -15.1506
8.6558 | -34.4027 | -15.1605
8.6619 | -34.403 | -15.1704
8.6681 | -34.4034 | -15.1805
8.6742 | -34.4038 | -15.1904
8.6803 | -34.4042 | -15.2002
8.6864 | -34.4046 | -15.2098
8.6925 | -34.4049 | -15.2196
8.6986 | -34.4053 | -15.2295
8.7047 | -34.4056 | -15.2395
8.7108 | -34.406 | -15.2494
8.7169 | -34.4064 | -15.2593
8.723 | -34.4067 | -15.2691
8.7291 | -34.4071 | -15.2788
8.7352 | -34.4075 | -15.2888
8.7413 | -34.4078 | -15.2988
8.7474 | -34.4082 | -15.3085
8.7535 | -34.4086 | -15.3184
8.7596 | -34.409 | -15.3283
8.7657 | -34.4093 | -15.3383
8.7718 | -34.4097 | -15.3482
8.7779 | -34.4101 | -15.3582
8.784 | -34.4104 | -15.3681
8.7901 | -34.4108 | -15.3779
8.7962 | -34.4112 | -15.3876
8.8023 | -34.4115 | -15.3974
8.8084 | -34.4119 | -15.4072
8.8146 | -34.4123 | -15.4171
8.8207 | -34.4126 | -15.4269
8.8268 | -34.413 | -15.4365
8.8329 | -34.4134 | -15.4464
8.839 | -34.4138 | -15.4562
8.8451 | -34.4141 | -15.466
8.8512 | -34.4145 | -15.4757
8.8573 | -34.4149 | -15.4856
8.8634 | -34.4153 | -15.4953
8.8695 | -34.4157 | -15.5053
8.8756 | -34.416 | -15.5152
8.8817 | -34.4164 | -15.5249
8.8878 | -34.4168 | -15.5349
8.8939 | -34.4171 | -15.5447
8.9 | -34.4175 | -15.5545
8.9061 | -34.4179 | -15.5643
8.9122 | -34.4183 | -15.5742
8.9183 | -34.4187 | -15.584
8.9244 | -34.419 | -15.594
8.9305 | -34.4195 | -15.6036
8.9366 | -34.4198 | -15.6134
8.9427 | -34.4202 | -15.6233
8.9488 | -34.4206 | -15.6331
8.955 | -34.421 | -15.6429
8.9611 | -34.4213 | -15.6529
8.9672 | -34.4218 | -15.6626
8.9733 | -34.4221 | -15.6723
8.9794 | -34.4225 | -15.682
8.9855 | -34.4229 | -15.6919
8.9916 | -34.4233 | -15.7018
8.9977 | -34.4237 | -15.7116
9.0038 | -34.4241 | -15.7213
9.0099 | -34.4245 | -15.7311
9.016 | -34.4248 | -15.741
9.0221 | -34.4252 | -15.7506
9.0282 | -34.4256 | -15.7602
9.0343 | -34.426 | -15.77
9.0404 | -34.4264 | -15.7797
9.0465 | -34.4268 | -15.7897
9.0526 | -34.4272 | -15.7994
9.0587 | -34.4276 | -15.8092
9.0648 | -34.428 | -15.8189
9.0709 | -34.4284 | -15.8286
9.077 | -34.4288 | -15.8385
9.0831 | -34.4292 | -15.8484
9.0892 | -34.4296 | -15.8582
9.0953 | -34.43 | -15.8681
9.1015 | -34.4303 | -15.8779
9.1076 | -34.4308 | -15.8876
9.1137 | -34.4311 | -15.8973
9.1198 | -34.4315 | -15.907
9.1259 | -34.4319 | -15.9168
9.132 | -34.4323 | -15.9268
9.1381 | -34.4327 | -15.9364
9.1442 | -34.4332 | -15.9459
9.1503 | -34.4335 | -15.9558
9.1564 | -34.4339 | -15.9655
9.1625 | -34.4343 | -15.9754
9.1686 | -34.4347 | -15.9851
9.1747 | -34.4351 | -15.9948
9.1808 | -34.4355 | -16.0046
9.1869 | -34.4359 | -16.0144
9.193 | -34.4363 | -16.0242
9.1991 | -34.4367 | -16.034
9.2052 | -34.4371 | -16.0437
9.2113 | -34.4375 | -16.0535
9.2174 | -34.438 | -16.0633
9.2235 | -34.4384 | -16.0731
9.2296 | -34.4388 | -16.0828
9.2357 | -34.4392 | -16.0927
9.2419 | -34.4396 | -16.1024
9.248 | -34.44 | -16.1122
9.2541 | -34.4404 | -16.122
9.2602 | -34.4408 | -16.1319
9.2663 | -34.4412 | -16.1414
9.2724 | -34.4416 | -16.1511
9.2785 | -34.442 | -16.1609
9.2846 | -34.4424 | -16.1705
9.2907 | -34.4429 | -16.1802
9.2968 | -34.4432 | -16.1902
9.3029 | -34.4437 | -16.2
9.309 | -34.4441 | -16.2098
9.3151 | -34.4445 | -16.2194
9.3212 | -34.4449 | -16.229
9.3273 | -34.4453 | -16.2387
9.3334 | -34.4457 | -16.2483
9.3395 | -34.4461 | -16.2579
9.3456 | -34.4465 | -16.2675
9.3517 | -34.4469 | -16.2773
9.3578 | -34.4473 | -16.2871
9.3639 | -34.4477 | -16.2966
9.37 | -34.4482 | -16.3064
9.3761 | -34.4486 | -16.3162
9.3822 | -34.449 | -16.326
9.3884 | -34.4494 | -16.3356
9.3945 | -34.4498 | -16.3454
9.4006 | -34.4503 | -16.3551
9.4067 | -34.4507 | -16.3648
9.4128 | -34.4511 | -16.3746
9.4189 | -34.4515 | -16.3845
9.425 | -34.4519 | -16.3941
9.4311 | -34.4523 | -16.4038
9.4372 | -34.4528 | -16.4135
9.4433 | -34.4532 | -16.4232
9.4494 | -34.4536 | -16.4328
9.4555 | -34.454 | -16.4426
9.4616 | -34.4544 | -16.4522
9.4677 | -34.4549 | -16.4619
9.4738 | -34.4553 | -16.4717
9.4799 | -34.4557 | -16.4814
9.486 | -34.4561 | -16.4912
9.4921 | -34.4565 | -16.5008
9.4982 | -34.457 | -16.5103
9.5043 | -34.4574 | -16.52
9.5104 | -34.4578 | -16.5296
9.5165 | -34.4582 | -16.5393
9.5226 | -34.4586 | -16.5492
9.5288 | -34.459 | -16.5587
9.5349 | -34.4594 | -16.5685
9.541 | -34.4599 | -16.578
9.5471 | -34.4603 | -16.5876
9.5532 | -34.4607 | -16.5975
9.5593 | -34.4611 | -16.6073
9.5654 | -34.4615 | -16.6169
9.5715 | -34.4619 | -16.6266
9.5776 | -34.4624 | -16.6364
9.5837 | -34.4628 | -16.646
9.5898 | -34.4632 | -16.6558
9.5959 | -34.4636 | -16.6656
9.602 | -34.464 | -16.6751
9.6081 | -34.4644 | -16.6847
9.6142 | -34.4649 | -16.6944
9.6203 | -34.4653 | -16.7041
9.6264 | -34.4657 | -16.7137
9.6325 | -34.4661 | -16.7235
9.6386 | -34.4665 | -16.7331
9.6447 | -34.4669 | -16.7428
9.6508 | -34.4673 | -16.7524
9.6569 | -34.4677 | -16.7621
9.663 | -34.4682 | -16.7719
9.6691 | -34.4686 | -16.7816
9.6753 | -34.469 | -16.7912
9.6814 | -34.4694 | -16.8007
9.6875 | -34.4699 | -16.8102
9.6936 | -34.4703 | -16.8196
9.6997 | -34.4707 | -16.8294
9.7058 | -34.4712 | -16.839
9.7119 | -34.4716 | -16.8488
9.718 | -34.472 | -16.8586
9.7241 | -34.4724 | -16.8681
9.7302 | -34.4728 | -16.8778
9.7363 | -34.4733 | -16.8874
9.7424 | -34.4737 | -16.8969
9.7485 | -34.4741 | -16.9066
9.7546 | -34.4745 | -16.9161
9.7607 | -34.475 | -16.9259
9.7668 | -34.4754 | -16.9356
9.7729 | -34.4758 | -16.9452
9.779 | -34.4763 | -16.9546
9.7851 | -34.4767 | -16.9644
9.7912 | -34.4771 | -16.974
9.7973 | -34.4776 | -16.9837
9.8034 | -34.478 | -16.9931
9.8095 | -34.4784 | -17.0029
9.8157 | -34.4789 | -17.0126
9.8218 | -34.4793 | -17.0224
9.8279 | -34.4797 | -17.0319
9.834 | -34.4801 | -17.0416
9.8401 | -34.4806 | -17.0513
9.8462 | -34.481 | -17.0607
9.8523 | -34.4814 | -17.0703
9.8584 | -34.4819 | -17.08
9.8645 | -34.4823 | -17.0897
9.8706 | -34.4827 | -17.0995
9.8767 | -34.4831 | -17.109
9.8828 | -34.4835 | -17.1185
9.8889 | -34.4839 | -17.1279
9.895 | -34.4844 | -17.1375
9.9011 | -34.4848 | -17.1471
9.9072 | -34.4852 | -17.1568
9.9133 | -34.4856 | -17.1664
9.9194 | -34.4861 | -17.1758
9.9255 | -34.4865 | -17.1853
9.9316 | -34.4869 | -17.1949
9.9377 | -34.4873 | -17.2044
9.9438 | -34.4877 | -17.2141
9.9499 | -34.4882 | -17.2239
9.956 | -34.4886 | -17.2335
9.9622 | -34.489 | -17.2432
9.9683 | -34.4894 | -17.2528
9.9744 | -34.4898 | -17.2625
9.9805 | -34.4903 | -17.2721
9.9866 | -34.4907 | -17.2816
9.9927 | -34.4911 | -17.2911
9.9988 | -34.4915 | -17.3008
10.0049 | -34.492 | -17.3104
10.011 | -34.4924 | -17.3201
10.0171 | -34.4928 | -17.3296
10.0232 | -34.4933 | -17.3391
10.0293 | -34.4937 | -17.3487
10.0354 | -34.4942 | -17.3584
10.0415 | -34.4946 | -17.368
10.0476 | -34.495 | -17.3777
10.0537 | -34.4954 | -17.3874
10.0598 | -34.4959 | -17.3969
10.0659 | -34.4963 | -17.4065
10.072 | -34.4967 | -17.416
10.0781 | -34.4971 | -17.4255
10.0842 | -34.4976 | -17.4351
10.0903 | -34.498 | -17.4446
10.0964 | -34.4984 | -17.454
10.1026 | -34.4988 | -17.4636
10.1087 | -34.4992 | -17.4731
10.1148 | -34.4996 | -17.4828
10.1209 | -34.5001 | -17.4923
10.127 | -34.5005 | -17.5016
10.1331 | -34.5009 | -17.5111
10.1392 | -34.5014 | -17.5205
10.1453 | -34.5018 | -17.53
10.1514 | -34.5022 | -17.5396
10.1575 | -34.5026 | -17.5492
10.1636 | -34.5031 | -17.5585
10.1697 | -34.5035 | -17.5682
10.1758 | -34.5039 | -17.5775
10.1819 | -34.5044 | -17.587
10.188 | -34.5047 | -17.5966
10.1941 | -34.5052 | -17.6061
10.2002 | -34.5056 | -17.6156
10.2063 | -34.506 | -17.6251
10.2124 | -34.5064 | -17.6348
10.2185 | -34.5069 | -17.6444
10.2246 | -34.5073 | -17.6539
10.2307 | -34.5077 | -17.6635
10.2368 | -34.5082 | -17.6726
10.2429 | -34.5086 | -17.6821
10.2491 | -34.509 | -17.6918
10.2552 | -34.5094 | -17.7013
10.2613 | -34.5098 | -17.7109
10.2674 | -34.5103 | -17.7204
10.2735 | -34.5107 | -17.7301
10.2796 | -34.5111 | -17.7395
10.2857 | -34.5116 | -17.7493
10.2918 | -34.512 | -17.7588
10.2979 | -34.5124 | -17.7684
10.304 | -34.5129 | -17.7777
10.3101 | -34.5133 | -17.7873
10.3162 | -34.5137 | -17.7968
10.3223 | -34.5141 | -17.8062
10.3284 | -34.5145 | -17.8157
10.3345 | -34.515 | -17.8251
10.3406 | -34.5154 | -17.8345
10.3467 | -34.5158 | -17.8439
10.3528 | -34.5162 | -17.8533
10.3589 | -34.5166 | -17.8628
10.365 | -34.5171 | -17.8724
10.3711 | -34.5175 | -17.882
10.3772 | -34.5179 | -17.8914
10.3833 | -34.5184 | -17.901
10.3895 | -34.5188 | -17.9104
10.3956 | -34.5192 | -17.92
10.4017 | -34.5197 | -17.9294
10.4078 | -34.5201 | -17.9389
10.4139 | -34.5205 | -17.9483
10.42 | -34.5209 | -17.9579
10.4261 | -34.5214 | -17.9673
10.4322 | -34.5218 | -17.9768
10.4383 | -34.5222 | -17.9863
10.4444 | -34.5226 | -17.9957
10.4505 | -34.523 | -18.0052
10.4566 | -34.5234 | -18.0145
10.4627 | -34.5238 | -18.0239
10.4688 | -34.5243 | -18.0331
10.4749 | -34.5247 | -18.0427
10.481 | -34.5251 | -18.0523
10.4871 | -34.5256 | -18.0616
10.4932 | -34.526 | -18.0711
10.4993 | -34.5264 | -18.0805
10.5054 | -34.5269 | -18.0899
10.5115 | -34.5273 | -18.0994
10.5176 | -34.5278 | -18.1087
10.5237 | -34.5282 | -18.1182
10.5298 | -34.5286 | -18.1276
10.536 | -34.529 | -18.1373
10.5421 | -34.5295 | -18.1467
10.5482 | -34.5299 | -18.1564
10.5543 | -34.5303 | -18.1658
10.5604 | -34.5308 | -18.175
10.5665 | -34.5312 | -18.1845
10.5726 | -34.5316 | -18.1938
10.5787 | -34.5321 | -18.2031
10.5848 | -34.5325 | -18.2122
10.5909 | -34.5329 | -18.2218
10.597 | -34.5334 | -18.2312
10.6031 | -34.5338 | -18.2409
10.6092 | -34.5342 | -18.2503
10.6153 | -34.5346 | -18.2598
10.6214 | -34.5351 | -18.269
10.6275 | -34.5355 | -18.2783
10.6336 | -34.5359 | -18.2878
10.6397 | -34.5364 | -18.2972
10.6458 | -34.5368 | -18.3066
10.6519 | -34.5372 | -18.3161
10.658 | -34.5376 | -18.3254
10.6641 | -34.5381 | -18.3351
10.6702 | -34.5385 | -18.3444
10.6764 | -34.5389 | -18.354
10.6825 | -34.5394 | -18.3634
10.6886 | -34.5398 | -18.3726
10.6947 | -34.5402 | -18.3819
10.7008 | -34.5407 | -18.3913
10.7069 | -34.5411 | -18.4009
10.713 | -34.5415 | -18.4102
10.7191 | -34.5419 | -18.4196
10.7252 | -34.5424 | -18.4287
10.7313 | -34.5428 | -18.4379
10.7374 | -34.5432 | -18.4473
10.7435 | -34.5436 | -18.4566
10.7496 | -34.5441 | -18.4658
10.7557 | -34.5445 | -18.4753
10.7618 | -34.5449 | -18.4845
10.7679 | -34.5454 | -18.4937
10.774 | -34.5458 | -18.5032
10.7801 | -34.5462 | -18.5127
10.7862 | -34.5466 | -18.5219
10.7923 | -34.547 | -18.5314
10.7984 | -34.5475 | -18.5406
10.8045 | -34.5479 | -18.5498
10.8106 | -34.5483 | -18.5592
10.8168 | -34.5487 | -18.5684
10.8229 | -34.5492 | -18.5777
10.829 | -34.5496 | -18.587
10.8351 | -34.5501 | -18.5959
10.8412 | -34.5505 | -18.6052
10.8473 | -34.5509 | -18.6144
10.8534 | -34.5514 | -18.6239
10.8595 | -34.5518 | -18.6332
10.8656 | -34.5523 | -18.6422
10.8717 | -34.5527 | -18.6514
10.8778 | -34.5531 | -18.6607
10.8839 | -34.5536 | -18.6701
10.89 | -34.554 | -18.6794
10.8961 | -34.5544 | -18.6886
10.9022 | -34.5549 | -18.6979
10.9083 | -34.5553 | -18.7071
10.9144 | -34.5557 | -18.7164
10.9205 | -34.5561 | -18.7258
10.9266 | -34.5566 | -18.7351
10.9327 | -34.557 | -18.7443
10.9388 | -34.5575 | -18.7535
10.9449 | -34.5579 | -18.7627
10.951 | -34.5583 | -18.7721
10.9571 | -34.5587 | -18.7815
10.9633 | -34.5591 | -18.7909
10.9694 | -34.5596 | -18.8004
10.9755 | -34.5601 | -18.8097
10.9816 | -34.5605 | -18.8191
10.9877 | -34.5609 | -18.8282
10.9938 | -34.5614 | -18.8376
10.9999 | -34.5618 | -18.8467
11.006 | -34.5622 | -18.8558
11.0121 | -34.5627 | -18.865
11.0182 | -34.5631 | -18.8743
11.0243 | -34.5635 | -18.8838
11.0304 | -34.5639 | -18.893
11.0365 | -34.5644 | -18.9022
11.0426 | -34.5648 | -18.9115
11.0487 | -34.5653 | -18.9209
11.0548 | -34.5657 | -18.9301
11.0609 | -34.5662 | -18.9394
11.067 | -34.5666 | -18.9487
11.0731 | -34.5671 | -18.9579
11.0792 | -34.5675 | -18.967
11.0853 | -34.5679 | -18.9764
11.0914 | -34.5684 | -18.9858
11.0975 | -34.5688 | -18.995
11.1037 | -34.5692 | -19.0041
11.1098 | -34.5697 | -19.0135
11.1159 | -34.5701 | -19.0225
11.122 | -34.5705 | -19.0316
11.1281 | -34.571 | -19.0409
11.1342 | -34.5714 | -19.0502
11.1403 | -34.5718 | -19.0596
11.1464 | -34.5723 | -19.069
11.1525 | -34.5727 | -19.078
11.1586 | -34.5731 | -19.0874
11.1647 | -34.5736 | -19.0966
11.1708 | -34.574 | -19.1059
11.1769 | -34.5745 | -19.1152
11.183 | -34.5749 | -19.1244
11.1891 | -34.5753 | -19.1337
11.1952 | -34.5758 | -19.1427
11.2013 | -34.5762 | -19.1517
11.2074 | -34.5767 | -19.1609
11.2135 | -34.5771 | -19.1702
11.2196 | -34.5776 | -19.1794
11.2257 | -34.578 | -19.1886
11.2318 | -34.5785 | -19.1979
11.2379 | -34.5789 | -19.2071
11.244 | -34.5794 | -19.2165
11.2502 | -34.5799 | -19.2257
11.2563 | -34.5803 | -19.2348
11.2624 | -34.5808 | -19.2441
11.2685 | -34.5812 | -19.2533
11.2746 | -34.5816 | -19.2625
11.2807 | -34.582 | -19.2719
11.2868 | -34.5825 | -19.2812
11.2929 | -34.583 | -19.2899
11.299 | -34.5834 | -19.2991
11.3051 | -34.5839 | -19.3087
11.3112 | -34.5843 | -19.318
11.3173 | -34.5847 | -19.3272
11.3234 | -34.5852 | -19.3363
11.3295 | -34.5857 | -19.3458
11.3356 | -34.5861 | -19.355
11.3417 | -34.5865 | -19.3642
11.3478 | -34.587 | -19.3734
11.3539 | -34.5874 | -19.3827
11.36 | -34.5878 | -19.3918
11.3661 | -34.5883 | -19.4009
11.3722 | -34.5888 | -19.41
11.3783 | -34.5892 | -19.4191
11.3844 | -34.5896 | -19.4283
11.3906 | -34.5901 | -19.4375
11.3967 | -34.5906 | -19.4465
11.4028 | -34.591 | -19.4558
11.4089 | -34.5915 | -19.4647
11.415 | -34.5919 | -19.474
11.4211 | -34.5923 | -19.4832
11.4272 | -34.5928 | -19.4924
11.4333 | -34.5932 | -19.5014
11.4394 | -34.5936 | -19.5106
11.4455 | -34.5941 | -19.5196
11.4516 | -34.5945 | -19.5288
11.4577 | -34.5949 | -19.5379
11.4638 | -34.5954 | -19.5469
11.4699 | -34.5959 | -19.5562
11.476 | -34.5963 | -19.5655
11.4821 | -34.5967 | -19.5746
11.4882 | -34.5972 | -19.5839
11.4943 | -34.5976 | -19.5932
11.5004 | -34.5981 | -19.6023
11.5065 | -34.5985 | -19.6116
11.5126 | -34.5989 | -19.6207
11.5187 | -34.5994 | -19.6299
11.5248 | -34.5998 | -19.6392
11.5309 | -34.6003 | -19.6483
11.5371 | -34.6007 | -19.6574
11.5432 | -34.6012 | -19.6665
11.5493 | -34.6016 | -19.6759
11.5554 | -34.6021 | -19.6851
11.5615 | -34.6025 | -19.6944
11.5676 | -34.603 | -19.7035
11.5737 | -34.6034 | -19.7125
11.5798 | -34.6039 | -19.7215
11.5859 | -34.6044 | -19.7306
11.592 | -34.6048 | -19.7396
11.5981 | -34.6052 | -19.7487
11.6042 | -34.6057 | -19.7579
11.6103 | -34.6061 | -19.767
11.6164 | -34.6065 | -19.776
11.6225 | -34.607 | -19.7852
11.6286 | -34.6074 | -19.7944
11.6347 | -34.6079 | -19.8035
11.6408 | -34.6083 | -19.8123
11.6469 | -34.6088 | -19.8214
11.653 | -34.6092 | -19.8305
11.6591 | -34.6097 | -19.8396
11.6652 | -34.6101 | -19.8486
11.6713 | -34.6106 | -19.8576
11.6775 | -34.611 | -19.8666
11.6836 | -34.6115 | -19.8757
11.6897 | -34.6119 | -19.8849
11.6958 | -34.6124 | -19.8941
11.7019 | -34.6129 | -19.9035
11.708 | -34.6133 | -19.9125
11.7141 | -34.6138 | -19.9215
11.7202 | -34.6142 | -19.9306
11.7263 | -34.6147 | -19.9398
11.7324 | -34.6151 | -19.9492
11.7385 | -34.6155 | -19.9584
11.7446 | -34.616 | -19.9674
11.7507 | -34.6164 | -19.9764
11.7568 | -34.6169 | -19.9855
11.7629 | -34.6173 | -19.9947
11.769 | -34.6177 | -20.0038
11.7751 | -34.6182 | -20.0129
11.7812 | -34.6187 | -20.0218
11.7873 | -34.6191 | -20.0308
11.7934 | -34.6196 | -20.0399
11.7995 | -34.62 | -20.0489
11.8056 | -34.6205 | -20.0579
11.8117 | -34.6209 | -20.067
11.8178 | -34.6214 | -20.076
11.824 | -34.6218 | -20.085
11.8301 | -34.6222 | -20.094
11.8362 | -34.6227 | -20.103
11.8423 | -34.6231 | -20.1121
11.8484 | -34.6236 | -20.1212
11.8545 | -34.6241 | -20.1302
11.8606 | -34.6245 | -20.1391
11.8667 | -34.625 | -20.1483
11.8728 | -34.6254 | -20.1573
11.8789 | -34.6258 | -20.1663
11.885 | -34.6263 | -20.1755
11.8911 | -34.6267 | -20.1845
11.8972 | -34.6272 | -20.1935
11.9033 | -34.6276 | -20.2024
11.9094 | -34.628 | -20.2114
11.9155 | -34.6285 | -20.2205
11.9216 | -34.629 | -20.2294
11.9277 | -34.6294 | -20.2384
11.9338 | -34.6298 | -20.2475
11.9399 | -34.6303 | -20.2567
11.946 | -34.6308 | -20.2657
11.9521 | -34.6312 | -20.2748
11.9582 | -34.6317 | -20.2838
11.9644 | -34.6321 | -20.2927
11.9705 | -34.6325 | -20.3017
11.9766 | -34.633 | -20.3107
11.9827 | -34.6335 | -20.3197
11.9888 | -34.634 | -20.3286
11.9949 | -34.6344 | -20.3373
12.001 | -34.6348 | -20.3464
12.0071 | -34.6353 | -20.3553
12.0132 | -34.6357 | -20.3642
12.0193 | -34.6362 | -20.3734
12.0254 | -34.6366 | -20.3823
12.0315 | -34.6371 | -20.3912
12.0376 | -34.6375 | -20.4002
12.0437 | -34.6379 | -20.4091
12.0498 | -34.6384 | -20.4182
12.0559 | -34.6388 | -20.4271
12.062 | -34.6393 | -20.4359
12.0681 | -34.6398 | -20.445
12.0742 | -34.6402 | -20.454
12.0803 | -34.6406 | -20.4631
12.0864 | -34.6411 | -20.4723
12.0925 | -34.6415 | -20.4813
12.0986 | -34.642 | -20.4902
12.1047 | -34.6425 | -20.4992
12.1109 | -34.6429 | -20.508
12.117 | -34.6434 | -20.5169
12.1231 | -34.6438 | -20.5259
12.1292 | -34.6443 | -20.5348
12.1353 | -34.6447 | -20.5437
12.1414 | -34.6452 | -20.5526
12.1475 | -34.6456 | -20.5617
12.1536 | -34.6461 | -20.5707
12.1597 | -34.6465 | -20.5796
12.1658 | -34.647 | -20.5885
12.1719 | -34.6474 | -20.5972
12.178 | -34.6479 | -20.6061
12.1841 | -34.6483 | -20.6151
12.1902 | -34.6488 | -20.6242
12.1963 | -34.6492 | -20.6333
12.2024 | -34.6497 | -20.6421
12.2085 | -34.6502 | -20.6511
12.2146 | -34.6506 | -20.66
12.2207 | -34.6511 | -20.6689
12.2268 | -34.6515 | -20.6779
12.2329 | -34.652 | -20.6867
12.239 | -34.6525 | -20.6954
12.2451 | -34.6529 | -20.7042
12.2513 | -34.6533 | -20.7133
12.2574 | -34.6538 | -20.722
12.2635 | -34.6542 | -20.7308
12.2696 | -34.6547 | -20.7396
12.2757 | -34.6552 | -20.7483
12.2818 | -34.6556 | -20.7572
12.2879 | -34.6561 | -20.7662
12.294 | -34.6565 | -20.775
12.3001 | -34.657 | -20.7839
12.3062 | -34.6575 | -20.7928
12.3123 | -34.6579 | -20.8013
12.3184 | -34.6583 | -20.8101
12.3245 | -34.6588 | -20.8188
12.3306 | -34.6592 | -20.8277
12.3367 | -34.6597 | -20.8365
12.3428 | -34.6602 | -20.8455
12.3489 | -34.6606 | -20.8541
12.355 | -34.6611 | -20.863
12.3611 | -34.6615 | -20.872
12.3672 | -34.662 | -20.8809
12.3733 | -34.6625 | -20.8899
12.3794 | -34.6629 | -20.8985
12.3855 | -34.6634 | -20.9073
12.3916 | -34.6638 | -20.9161
12.3978 | -34.6643 | -20.9251
12.4039 | -34.6648 | -20.934
12.41 | -34.6652 | -20.9428
12.4161 | -34.6657 | -20.9514
12.4222 | -34.6662 | -20.9602
12.4283 | -34.6666 | -20.969
12.4344 | -34.6671 | -20.978
12.4405 | -34.6675 | -20.9869
12.4466 | -34.668 | -20.9957
12.4527 | -34.6684 | -21.0044
12.4588 | -34.6689 | -21.0132
12.4649 | -34.6693 | -21.0222
12.471 | -34.6698 | -21.0311
12.4771 | -34.6702 | -21.0399
12.4832 | -34.6707 | -21.0487
12.4893 | -34.6711 | -21.0574
12.4954 | -34.6716 | -21.066
12.5015 | -34.672 | -21.0747
12.5076 | -34.6725 | -21.0836
12.5137 | -34.673 | -21.0923
12.5198 | -34.6734 | -21.101
12.5259 | -34.6739 | -21.1099
12.532 | -34.6743 | -21.1186
12.5382 | -34.6748 | -21.1274
12.5443 | -34.6752 | -21.1361
12.5504 | -34.6757 | -21.145
12.5565 | -34.6761 | -21.1535
12.5626 | -34.6766 | -21.1624
12.5687 | -34.677 | -21.1713
12.5748 | -34.6775 | -21.1801
12.5809 | -34.6779 | -21.189
12.587 | -34.6784 | -21.1978
12.5931 | -34.6789 | -21.2064
12.5992 | -34.6794 | -21.2152
12.6053 | -34.6798 | -21.224
12.6114 | -34.6803 | -21.2326
12.6175 | -34.6807 | -21.2415
12.6236 | -34.6812 | -21.2503
12.6297 | -34.6816 | -21.259
12.6358 | -34.6821 | -21.2677
12.6419 | -34.6825 | -21.2765
12.648 | -34.683 | -21.2853
12.6541 | -34.6834 | -21.2941
12.6602 | -34.6839 | -21.3026
12.6663 | -34.6843 | -21.3114
12.6724 | -34.6848 | -21.32
12.6785 | -34.6852 | -21.3287
12.6847 | -34.6857 | -21.3373
12.6908 | -34.6862 | -21.3461
12.6969 | -34.6867 | -21.3547
12.703 | -34.6871 | -21.3635
12.7091 | -34.6875 | -21.3722
12.7152 | -34.688 | -21.3809
12.7213 | -34.6884 | -21.3896
12.7274 | -34.6889 | -21.3983
12.7335 | -34.6894 | -21.4068
12.7396 | -34.6898 | -21.4155
12.7457 | -34.6903 | -21.4243
12.7518 | -34.6907 | -21.433
12.7579 | -34.6912 | -21.4418
12.764 | -34.6916 | -21.4506
12.7701 | -34.6921 | -21.4592
12.7762 | -34.6925 | -21.4678
12.7823 | -34.693 | -21.4766
12.7884 | -34.6935 | -21.4854
12.7945 | -34.6939 | -21.4943
12.8006 | -34.6944 | -21.5028
12.8067 | -34.6949 | -21.5114
12.8128 | -34.6953 | -21.52
12.8189 | -34.6958 | -21.5287
12.8251 | -34.6962 | -21.5374
12.8312 | -34.6967 | -21.5462
12.8373 | -34.6971 | -21.5547
12.8434 | -34.6976 | -21.5635
12.8495 | -34.698 | -21.5721
12.8556 | -34.6985 | -21.5807
12.8617 | -34.6989 | -21.5895
12.8678 | -34.6994 | -21.5983
12.8739 | -34.6998 | -21.6069
12.88 | -34.7003 | -21.6155
12.8861 | -34.7007 | -21.6242
12.8922 | -34.7012 | -21.6329
12.8983 | -34.7016 | -21.6416
12.9044 | -34.7021 | -21.6501
12.9105 | -34.7026 | -21.6587
12.9166 | -34.703 | -21.6674
12.9227 | -34.7035 | -21.6761
12.9288 | -34.7039 | -21.6849
12.9349 | -34.7044 | -21.6935
12.941 | -34.7049 | -21.7021
12.9471 | -34.7053 | -21.7107
12.9532 | -34.7058 | -21.7192
12.9593 | -34.7063 | -21.7279
12.9654 | -34.7067 | -21.7367
12.9716 | -34.7072 | -21.7455
12.9777 | -34.7076 | -21.7541
12.9838 | -34.7081 | -21.7626
12.9899 | -34.7086 | -21.7713
12.996 | -34.709 | -21.78
13.0021 | -34.7095 | -21.7887
13.0082 | -34.7099 | -21.7973
13.0143 | -34.7104 | -21.8057
13.0204 | -34.7109 | -21.8142
13.0265 | -34.7113 | -21.8227
13.0326 | -34.7118 | -21.8314
13.0387 | -34.7122 | -21.8402
13.0448 | -34.7127 | -21.8488
13.0509 | -34.7132 | -21.8573
13.057 | -34.7136 | -21.8658
13.0631 | -34.7141 | -21.8744
13.0692 | -34.7146 | -21.883
13.0753 | -34.715 | -21.8917
13.0814 | -34.7155 | -21.9005
13.0875 | -34.7159 | -21.9089
13.0936 | -34.7164 | -21.9175
13.0997 | -34.7169 | -21.9261
13.1058 | -34.7173 | -21.9347
13.112 | -34.7178 | -21.9433
13.1181 | -34.7183 | -21.9519
13.1242 | -34.7187 | -21.9606
13.1303 | -34.7192 | -21.9692
13.1364 | -34.7196 | -21.9778
13.1425 | -34.72 | -21.9864
13.1486 | -34.7205 | -21.9949
13.1547 | -34.721 | -22.0033
13.1608 | -34.7214 | -22.012
13.1669 | -34.7219 | -22.0205
13.173 | -34.7224 | -22.0291
13.1791 | -34.7228 | -22.0376
13.1852 | -34.7233 | -22.0462
13.1913 | -34.7237 | -22.0545
13.1974 | -34.7242 | -22.0632
13.2035 | -34.7246 | -22.0718
13.2096 | -34.7251 | -22.0803
13.2157 | -34.7256 | -22.0888
13.2218 | -34.726 | -22.0974
13.2279 | -34.7265 | -22.1059
13.234 | -34.727 | -22.1144
13.2401 | -34.7274 | -22.123
13.2462 | -34.7279 | -22.1315
13.2524 | -34.7283 | -22.1402
13.2585 | -34.7288 | -22.1486
13.2646 | -34.7292 | -22.1572
13.2707 | -34.7297 | -22.1658
13.2768 | -34.7302 | -22.1743
13.2829 | -34.7306 | -22.1831
13.289 | -34.7311 | -22.1916
13.2951 | -34.7316 | -22.2
13.3012 | -34.732 | -22.2085
13.3073 | -34.7325 | -22.2171
13.3134 | -34.7329 | -22.2258
13.3195 | -34.7334 | -22.2344
13.3256 | -34.7338 | -22.2427
13.3317 | -34.7343 | -22.251
13.3378 | -34.7348 | -22.2596
13.3439 | -34.7353 | -22.2682
13.35 | -34.7357 | -22.2768
13.3561 | -34.7362 | -22.2853
13.3622 | -34.7366 | -22.2937
13.3683 | -34.7371 | -22.3021
13.3744 | -34.7375 | -22.3105
13.3805 | -34.738 | -22.319
13.3866 | -34.7385 | -22.3274
13.3927 | -34.7389 | -22.3359
13.3989 | -34.7394 | -22.3444
13.405 | -34.7398 | -22.3528
13.4111 | -34.7403 | -22.3613
13.4172 | -34.7407 | -22.3698
13.4233 | -34.7412 | -22.3784
13.4294 | -34.7416 | -22.3871
13.4355 | -34.7421 | -22.3954
13.4416 | -34.7426 | -22.4038
13.4477 | -34.743 | -22.4123
13.4538 | -34.7435 | -22.4209
13.4599 | -34.744 | -22.4294
13.466 | -34.7444 | -22.438
13.4721 | -34.7449 | -22.4464
13.4782 | -34.7453 | -22.4548
13.4843 | -34.7458 | -22.4634
13.4904 | -34.7463 | -22.4719
13.4965 | -34.7467 | -22.4805
13.5026 | -34.7472 | -22.489
13.5087 | -34.7476 | -22.4974
13.5148 | -34.7481 | -22.5059
13.5209 | -34.7485 | -22.5146
13.527 | -34.7489 | -22.5231
13.5331 | -34.7494 | -22.5315
13.5393 | -34.7499 | -22.5399
13.5454 | -34.7504 | -22.5483
13.5515 | -34.7508 | -22.5569
13.5576 | -34.7513 | -22.5653
13.5637 | -34.7518 | -22.5738
13.5698 | -34.7522 | -22.5823
13.5759 | -34.7527 | -22.5906
13.582 | -34.7531 | -22.5989
13.5881 | -34.7536 | -22.6073
13.5942 | -34.754 | -22.616
13.6003 | -34.7545 | -22.6245
13.6064 | -34.7549 | -22.6331
13.6125 | -34.7554 | -22.6415
13.6186 | -34.7559 | -22.6499
13.6247 | -34.7563 | -22.6585
13.6308 | -34.7568 | -22.667
13.6369 | -34.7572 | -22.6755
13.643 | -34.7577 | -22.6841
13.6491 | -34.7582 | -22.6924
13.6552 | -34.7586 | -22.7009
13.6613 | -34.7591 | -22.7093
13.6674 | -34.7595 | -22.7178
13.6735 | -34.76 | -22.7262
13.6796 | -34.7604 | -22.7347
13.6858 | -34.7609 | -22.7431
13.6919 | -34.7614 | -22.7515
13.698 | -34.7618 | -22.76
13.7041 | -34.7623 | -22.7685
13.7102 | -34.7628 | -22.7769
13.7163 | -34.7632 | -22.7853
13.7224 | -34.7637 | -22.7938
13.7285 | -34.7641 | -22.8022
13.7346 | -34.7646 | -22.8108
13.7407 | -34.765 | -22.8192
13.7468 | -34.7655 | -22.8277
13.7529 | -34.766 | -22.8359
13.759 | -34.7664 | -22.8445
13.7651 | -34.7669 | -22.8529
13.7712 | -34.7673 | -22.8614
13.7773 | -34.7678 | -22.8699
13.7834 | -34.7682 | -22.8784
13.7895 | -34.7687 | -22.8867
13.7956 | -34.7692 | -22.8952
13.8017 | -34.7696 | -22.9036
13.8078 | -34.7701 | -22.9121
13.8139 | -34.7705 | -22.9205
13.82 | -34.771 | -22.9288
13.8262 | -34.7714 | -22.9372
13.8323 | -34.7719 | -22.9458
13.8384 | -34.7724 | -22.9543
13.8445 | -34.7728 | -22.9627
13.8506 | -34.7733 | -22.9712
13.8567 | -34.7738 | -22.9796
13.8628 | -34.7742 | -22.9881
13.8689 | -34.7747 | -22.9966
13.875 | -34.7751 | -23.0051
13.8811 | -34.7756 | -23.0135
13.8872 | -34.776 | -23.0219
13.8933 | -34.7765 | -23.0301
13.8994 | -34.777 | -23.0386
13.9055 | -34.7774 | -23.0472
13.9116 | -34.7779 | -23.0555
13.9177 | -34.7783 | -23.0641
13.9238 | -34.7788 | -23.0725
13.9299 | -34.7792 | -23.0808
13.936 | -34.7797 | -23.0893
13.9421 | -34.7801 | -23.0977
13.9482 | -34.7805 | -23.1061
13.9543 | -34.781 | -23.1146
13.9604 | -34.7814 | -23.123
13.9665 | -34.7819 | -23.1314
13.9727 | -34.7824 | -23.14
13.9788 | -34.7829 | -23.1484
13.9849 | -34.7833 | -23.1568
13.991 | -34.7838 | -23.1652
13.9971 | -34.7842 | -23.1735
14.0032 | -34.7847 | -23.182
14.0093 | -34.7852 | -23.1904
14.0154 | -34.7856 | -23.1987
14.0215 | -34.7861 | -23.2071
14.0276 | -34.7865 | -23.2154
14.0337 | -34.787 | -23.2237
14.0398 | -34.7875 | -23.2321
14.0459 | -34.7879 | -23.2405
14.052 | -34.7884 | -23.249
14.0581 | -34.7888 | -23.2575
14.0642 | -34.7893 | -23.2658
14.0703 | -34.7898 | -23.2742
14.0764 | -34.7902 | -23.2825
14.0825 | -34.7907 | -23.2908
14.0886 | -34.7912 | -23.2991
14.0947 | -34.7916 | -23.3076
14.1008 | -34.7921 | -23.3161
14.1069 | -34.7925 | -23.3244
14.1131 | -34.793 | -23.3329
14.1192 | -34.7934 | -23.3413
14.1253 | -34.7939 | -23.3495
14.1314 | -34.7943 | -23.3579
14.1375 | -34.7948 | -23.3662
14.1436 | -34.7953 | -23.3745
14.1497 | -34.7958 | -23.3828
14.1558 | -34.7962 | -23.3911
14.1619 | -34.7967 | -23.3993
14.168 | -34.7971 | -23.4077
14.1741 | -34.7976 | -23.4159
14.1802 | -34.7981 | -23.4243
14.1863 | -34.7986 | -23.4328
14.1924 | -34.799 | -23.4412
14.1985 | -34.7995 | -23.4495
14.2046 | -34.8 | -23.4579
14.2107 | -34.8004 | -23.4661
14.2168 | -34.8009 | -23.4744
14.2229 | -34.8013 | -23.4828
14.229 | -34.8018 | -23.4911
14.2351 | -34.8023 | -23.4993
14.2412 | -34.8027 | -23.5077
14.2473 | -34.8032 | -23.5159
14.2534 | -34.8037 | -23.5241
14.2596 | -34.8041 | -23.5324
14.2657 | -34.8046 | -23.5407
14.2718 | -34.805 | -23.549
14.2779 | -34.8055 | -23.5574
14.284 | -34.806 | -23.5658
14.2901 | -34.8065 | -23.5742
14.2962 | -34.8069 | -23.5826
14.3023 | -34.8074 | -23.591
14.3084 | -34.8079 | -23.5994
14.3145 | -34.8083 | -23.6077
14.3206 | -34.8088 | -23.6162
14.3267 | -34.8092 | -23.6246
14.3328 | -34.8097 | -23.6329
14.3389 | -34.8102 | -23.6412
14.345 | -34.8106 | -23.6495
14.3511 | -34.8111 | -23.6578
14.3572 | -34.8116 | -23.6663
14.3633 | -34.812 | -23.6747
14.3694 | -34.8125 | -23.6829
14.3755 | -34.813 | -23.6913
14.3816 | -34.8135 | -23.6996
14.3877 | -34.814 | -23.7079
14.3938 | -34.8144 | -23.7165
14.4 | -34.8149 | -23.7247
14.4061 | -34.8153 | -23.7331
14.4122 | -34.8158 | -23.7414
14.4183 | -34.8162 | -23.7497
14.4244 | -34.8167 | -23.758
14.4305 | -34.8172 | -23.7663
14.4366 | -34.8177 | -23.7746
14.4427 | -34.8181 | -23.7831
14.4488 | -34.8186 | -23.7915
14.4549 | -34.8191 | -23.7997
14.461 | -34.8195 | -23.8079
14.4671 | -34.82 | -23.8164
14.4732 | -34.8204 | -23.8247
14.4793 | -34.8209 | -23.8331
14.4854 | -34.8214 | -23.8415
14.4915 | -34.8218 | -23.8496
14.4976 | -34.8223 | -23.8578
14.5037 | -34.8228 | -23.8661
14.5098 | -34.8232 | -23.8744
14.5159 | -34.8237 | -23.8827
14.522 | -34.8241 | -23.8909
14.5281 | -34.8246 | -23.8992
14.5342 | -34.8251 | -23.9077
14.5403 | -34.8256 | -23.9159
14.5465 | -34.826 | -23.9243
14.5526 | -34.8265 | -23.9327
14.5587 | -34.827 | -23.941
14.5648 | -34.8274 | -23.9493
14.5709 | -34.8279 | -23.9575
14.577 | -34.8284 | -23.9658
14.5831 | -34.8288 | -23.9742
14.5892 | -34.8293 | -23.9825
14.5953 | -34.8298 | -23.9905
14.6014 | -34.8302 | -23.9988
14.6075 | -34.8307 | -24.007
14.6136 | -34.8311 | -24.0153
14.6197 | -34.8316 | -24.0237
14.6258 | -34.832 | -24.032
14.6319 | -34.8325 | -24.0402
14.638 | -34.8329 | -24.0485
14.6441 | -34.8334 | -24.0568
14.6502 | -34.8339 | -24.0652
14.6563 | -34.8344 | -24.0735
14.6624 | -34.8348 | -24.0818
14.6685 | -34.8353 | -24.09
14.6746 | -34.8358 | -24.0985
14.6807 | -34.8362 | -24.1065
14.6869 | -34.8367 | -24.1148
14.693 | -34.8372 | -24.1231
14.6991 | -34.8376 | -24.1313
14.7052 | -34.8381 | -24.1394
14.7113 | -34.8386 | -24.1478
14.7174 | -34.839 | -24.1562
14.7235 | -34.8395 | -24.1643
14.7296 | -34.84 | -24.1727
14.7357 | -34.8405 | -24.1809
14.7418 | -34.8409 | -24.1893
14.7479 | -34.8414 | -24.1975
14.754 | -34.8419 | -24.2059
14.7601 | -34.8423 | -24.2141
14.7662 | -34.8428 | -24.2223
14.7723 | -34.8433 | -24.2304
14.7784 | -34.8437 | -24.2386
14.7845 | -34.8442 | -24.2469
14.7906 | -34.8447 | -24.2554
14.7967 | -34.8452 | -24.2636
14.8028 | -34.8456 | -24.2718
14.8089 | -34.8461 | -24.2799
14.815 | -34.8466 | -24.2882
14.8211 | -34.847 | -24.2964
14.8272 | -34.8475 | -24.3048
14.8334 | -34.848 | -24.313
14.8395 | -34.8485 | -24.3212
14.8456 | -34.8489 | -24.3293
14.8517 | -34.8494 | -24.3376
14.8578 | -34.8499 | -24.346
14.8639 | -34.8504 | -24.3543
14.87 | -34.8508 | -24.3626
14.8761 | -34.8513 | -24.3709
14.8822 | -34.8518 | -24.3792
14.8883 | -34.8522 | -24.3876
14.8944 | -34.8527 | -24.3957
14.9005 | -34.8532 | -24.4038
14.9066 | -34.8537 | -24.412
14.9127 | -34.8542 | -24.42
14.9188 | -34.8546 | -24.4282
14.9249 | -34.8551 | -24.4365
14.931 | -34.8556 | -24.4448
14.9371 | -34.856 | -24.453
14.9432 | -34.8565 | -24.4612
14.9493 | -34.857 | -24.4694
14.9554 | -34.8575 | -24.4776
14.9615 | -34.8579 | -24.4859
14.9676 | -34.8584 | -24.4941
14.9738 | -34.8588 | -24.5022
14.9799 | -34.8593 | -24.5105
14.986 | -34.8598 | -24.5186
14.9921 | -34.8603 | -24.5267
14.9982 | -34.8607 | -24.5348
15.0043 | -34.8612 | -24.5429
15.0104 | -34.8616 | -24.5513
15.0165 | -34.8621 | -24.5596
15.0226 | -34.8626 | -24.5677
15.0287 | -34.8631 | -24.576
15.0348 | -34.8635 | -24.5843
15.0409 | -34.864 | -24.5927
15.047 | -34.8645 | -24.601
15.0531 | -34.865 | -24.609
15.0592 | -34.8655 | -24.6173
15.0653 | -34.8659 | -24.6256
15.0714 | -34.8664 | -24.6336
15.0775 | -34.8669 | -24.6418
15.0836 | -34.8674 | -24.6501
15.0897 | -34.8679 | -24.658
15.0958 | -34.8684 | -24.6663
15.1019 | -34.8688 | -24.6746
15.108 | -34.8693 | -24.6827
15.1141 | -34.8698 | -24.6909
15.1203 | -34.8703 | -24.6992
15.1264 | -34.8707 | -24.7074
15.1325 | -34.8712 | -24.7157
15.1386 | -34.8717 | -24.7239
15.1447 | -34.8722 | -24.7322
15.1508 | -34.8727 | -24.7405
15.1569 | -34.8732 | -24.7485
15.163 | -34.8737 | -24.7566
15.1691 | -34.8742 | -24.7647
15.1752 | -34.8746 | -24.7727
15.1813 | -34.8751 | -24.7809
15.1874 | -34.8756 | -24.7891
15.1935 | -34.8761 | -24.7972
15.1996 | -34.8765 | -24.8054
15.2057 | -34.877 | -24.8136
15.2118 | -34.8775 | -24.8218
15.2179 | -34.878 | -24.83
15.224 | -34.8785 | -24.8383
15.2301 | -34.879 | -24.8463
15.2362 | -34.8794 | -24.8544
15.2423 | -34.8799 | -24.8628
15.2484 | -34.8804 | -24.871
15.2545 | -34.8809 | -24.8792
15.2607 | -34.8814 | -24.8873
15.2668 | -34.8819 | -24.8953
15.2729 | -34.8824 | -24.9035
15.279 | -34.8829 | -24.9118
15.2851 | -34.8834 | -24.9198
15.2912 | -34.8838 | -24.928
15.2973 | -34.8843 | -24.9363
15.3034 | -34.8848 | -24.9443
15.3095 | -34.8853 | -24.9526
15.3156 | -34.8858 | -24.9606
15.3217 | -34.8863 | -24.9687
15.3278 | -34.8868 | -24.977
15.3339 | -34.8873 | -24.985
15.34 | -34.8878 | -24.9931
15.3461 | -34.8882 | -25.0011
15.3522 | -34.8887 | -25.0093
15.3583 | -34.8892 | -25.0176
15.3644 | -34.8897 | -25.0258
15.3705 | -34.8902 | -25.0337
15.3766 | -34.8907 | -25.0419
15.3827 | -34.8911 | -25.0501
15.3888 | -34.8916 | -25.0583
15.3949 | -34.8921 | -25.0664
15.401 | -34.8926 | -25.0747
15.4072 | -34.8931 | -25.0826
15.4133 | -34.8935 | -25.0907
15.4194 | -34.894 | -25.0988
15.4255 | -34.8945 | -25.1069
15.4316 | -34.895 | -25.115
15.4377 | -34.8955 | -25.1233
15.4438 | -34.896 | -25.1314
15.4499 | -34.8965 | -25.1394
15.456 | -34.8969 | -25.1475
15.4621 | -34.8974 | -25.1556
15.4682 | -34.8979 | -25.1639
15.4743 | -34.8984 | -25.1718
15.4804 | -34.8989 | -25.1799
15.4865 | -34.8994 | -25.1881
15.4926 | -34.8999 | -25.1963
15.4987 | -34.9004 | -25.2044
15.5048 | -34.9008 | -25.2126
15.5109 | -34.9013 | -25.2205
15.517 | -34.9018 | -25.2287
15.5231 | -34.9023 | -25.2367
15.5292 | -34.9028 | -25.2447
15.5353 | -34.9033 | -25.2529
15.5414 | -34.9038 | -25.2609
15.5476 | -34.9043 | -25.2688
15.5537 | -34.9048 | -25.2768
15.5598 | -34.9053 | -25.2851
15.5659 | -34.9058 | -25.2932
15.572 | -34.9063 | -25.3012
15.5781 | -34.9068 | -25.3091
15.5842 | -34.9073 | -25.317
15.5903 | -34.9078 | -25.3249
15.5964 | -34.9083 | -25.3331
15.6025 | -34.9088 | -25.341
15.6086 | -34.9092 | -25.3492
15.6147 | -34.9097 | -25.3573
15.6208 | -34.9102 | -25.3651
15.6269 | -34.9107 | -25.3732
15.633 | -34.9112 | -25.3815
15.6391 | -34.9117 | -25.3896
15.6452 | -34.9122 | -25.3977
15.6513 | -34.9127 | -25.4056
15.6574 | -34.9132 | -25.4138
15.6635 | -34.9137 | -25.4218
15.6696 | -34.9142 | -25.4298
15.6757 | -34.9147 | -25.4379
15.6818 | -34.9152 | -25.446
15.688 | -34.9157 | -25.4539
15.6941 | -34.9162 | -25.4619
15.7002 | -34.9166 | -25.4699
15.7063 | -34.9171 | -25.4779
15.7124 | -34.9176 | -25.486
15.7185 | -34.918 | -25.494
15.7246 | -34.9186 | -25.502
15.7307 | -34.9191 | -25.51
15.7368 | -34.9196 | -25.518
15.7429 | -34.9201 | -25.5261
15.749 | -34.9206 | -25.5342
15.7551 | -34.9211 | -25.5421
15.7612 | -34.9216 | -25.55
15.7673 | -34.9221 | -25.5581
15.7734 | -34.9226 | -25.5661
15.7795 | -34.923 | -25.5742
15.7856 | -34.9235 | -25.5821
15.7917 | -34.9241 | -25.59
15.7978 | -34.9245 | -25.5981
15.8039 | -34.925 | -25.6061
15.81 | -34.9255 | -25.6141
15.8161 | -34.926 | -25.6221
15.8222 | -34.9265 | -25.63
15.8283 | -34.927 | -25.6378
15.8345 | -34.9275 | -25.6458
15.8406 | -34.928 | -25.6537
15.8467 | -34.9285 | -25.6618
15.8528 | -34.929 | -25.6699
15.8589 | -34.9296 | -25.6779
15.865 | -34.9301 | -25.6856
15.8711 | -34.9306 | -25.6936
15.8772 | -34.931 | -25.7014
15.8833 | -34.9315 | -25.7095
15.8894 | -34.9321 | -25.7176
15.8955 | -34.9326 | -25.7255
15.9016 | -34.9331 | -25.7333
15.9077 | -34.9336 | -25.7412
15.9138 | -34.934 | -25.7494
15.9199 | -34.9345 | -25.7575
15.926 | -34.935 | -25.7655
15.9321 | -34.9355 | -25.7734
15.9382 | -34.936 | -25.7814
15.9443 | -34.9365 | -25.7894
15.9504 | -34.937 | -25.7974
15.9565 | -34.9375 | -25.8053
15.9626 | -34.938 | -25.8132
15.9687 | -34.9385 | -25.8212
15.9749 | -34.939 | -25.8292
15.981 | -34.9395 | -25.8372
15.9871 | -34.9399 | -25.8452
15.9932 | -34.9404 | -25.8531
15.9993 | -34.9409 | -25.8612
16.0054 | -34.9414 | -25.8689
16.0115 | -34.942 | -25.8769
16.0176 | -34.9425 | -25.8849
16.0237 | -34.943 | -25.8929
16.0298 | -34.9435 | -25.9009
16.0359 | -34.944 | -25.9088
16.042 | -34.9445 | -25.9167
16.0481 | -34.9449 | -25.9249
16.0542 | -34.9454 | -25.9329
16.0603 | -34.9459 | -25.9409
16.0664 | -34.9464 | -25.9487
16.0725 | -34.9469 | -25.9563
16.0786 | -34.9475 | -25.9642
16.0847 | -34.948 | -25.9722
16.0908 | -34.9484 | -25.9802
16.0969 | -34.9489 | -25.9881
16.103 | -34.9494 | -25.996
16.1091 | -34.9499 | -26.0037
16.1152 | -34.9504 | -26.0117
16.1214 | -34.9509 | -26.0195
16.1275 | -34.9514 | -26.0275
16.1336 | -34.9519 | -26.0353
16.1397 | -34.9524 | -26.0433
16.1458 | -34.9529 | -26.0511
16.1519 | -34.9535 | -26.0591
16.158 | -34.954 | -26.0671
16.1641 | -34.9545 | -26.0751
16.1702 | -34.9549 | -26.0833
16.1763 | -34.9555 | -26.0912
16.1824 | -34.956 | -26.0989
16.1885 | -34.9564 | -26.1066
16.1946 | -34.9569 | -26.1147
16.2007 | -34.9574 | -26.1227
16.2068 | -34.9579 | -26.1307
16.2129 | -34.9584 | -26.1388
16.219 | -34.959 | -26.1463
16.2251 | -34.9595 | -26.1542
16.2312 | -34.96 | -26.162
16.2373 | -34.9605 | -26.1699
16.2434 | -34.961 | -26.1777
16.2495 | -34.9615 | -26.1854
16.2556 | -34.962 | -26.1934
16.2618 | -34.9625 | -26.2012
16.2679 | -34.963 | -26.2092
16.274 | -34.9635 | -26.2172
16.2801 | -34.964 | -26.2251
16.2862 | -34.9645 | -26.2329
16.2923 | -34.965 | -26.2408
16.2984 | -34.9655 | -26.2487
16.3045 | -34.966 | -26.2565
16.3106 | -34.9665 | -26.2645
16.3167 | -34.967 | -26.2723
16.3228 | -34.9675 | -26.2799
16.3289 | -34.968 | -26.2877
16.335 | -34.9685 | -26.2955
16.3411 | -34.969 | -26.3032
16.3472 | -34.9695 | -26.311
16.3533 | -34.97 | -26.3187
16.3594 | -34.9705 | -26.3264
16.3655 | -34.971 | -26.3342
16.3716 | -34.9715 | -26.3419
16.3777 | -34.972 | -26.3498
16.3838 | -34.9725 | -26.3577
16.3899 | -34.973 | -26.3656
16.396 | -34.9735 | -26.3733
16.4021 | -34.974 | -26.3811
16.4083 | -34.9745 | -26.3889
16.4144 | -34.975 | -26.3966
16.4205 | -34.9755 | -26.4044
16.4266 | -34.976 | -26.412
16.4327 | -34.9765 | -26.4198
16.4388 | -34.977 | -26.4277
16.4449 | -34.9775 | -26.4355
16.451 | -34.978 | -26.4434
16.4571 | -34.9785 | -26.4513
16.4632 | -34.979 | -26.4589
16.4693 | -34.9795 | -26.4668
16.4754 | -34.98 | -26.4748
16.4815 | -34.9805 | -26.4826
16.4876 | -34.981 | -26.4902
16.4937 | -34.9815 | -26.498
16.4998 | -34.982 | -26.5056
16.5059 | -34.9825 | -26.5134
16.512 | -34.983 | -26.5212
16.5181 | -34.9836 | -26.5291
16.5242 | -34.984 | -26.5369
16.5303 | -34.9845 | -26.5446
16.5364 | -34.985 | -26.5524
16.5425 | -34.9855 | -26.5603
16.5487 | -34.986 | -26.5682
16.5548 | -34.9865 | -26.576
16.5609 | -34.987 | -26.5838
16.567 | -34.9876 | -26.5913
16.5731 | -34.9881 | -26.5993
16.5792 | -34.9886 | -26.6072
16.5853 | -34.9891 | -26.6151
16.5914 | -34.9896 | -26.6228
16.5975 | -34.9901 | -26.6307
16.6036 | -34.9906 | -26.6384
16.6097 | -34.9911 | -26.6462
16.6158 | -34.9916 | -26.6539
16.6219 | -34.9921 | -26.6616
16.628 | -34.9926 | -26.6694
16.6341 | -34.9931 | -26.6773
16.6402 | -34.9936 | -26.6848
16.6463 | -34.994 | -26.6927
16.6524 | -34.9945 | -26.7003
16.6585 | -34.9951 | -26.7081
16.6646 | -34.9956 | -26.7159
16.6707 | -34.9961 | -26.7234
16.6768 | -34.9966 | -26.7311
16.6829 | -34.9971 | -26.7386
16.689 | -34.9976 | -26.7464
16.6952 | -34.9981 | -26.7542
16.7013 | -34.9986 | -26.7619
16.7074 | -34.9991 | -26.7693
16.7135 | -34.9996 | -26.7771
16.7196 | -35.0001 | -26.7848
16.7257 | -35.0007 | -26.7925
16.7318 | -35.0012 | -26.8002
16.7379 | -35.0017 | -26.8079
16.744 | -35.0022 | -26.8154
16.7501 | -35.0027 | -26.8231
16.7562 | -35.0032 | -26.8309
16.7623 | -35.0037 | -26.8386
16.7684 | -35.0042 | -26.8464
16.7745 | -35.0047 | -26.8542
16.7806 | -35.0052 | -26.8616
16.7867 | -35.0057 | -26.8693
16.7928 | -35.0062 | -26.8772
16.7989 | -35.0067 | -26.8848
16.805 | -35.0072 | -26.8927
16.8111 | -35.0077 | -26.9001
16.8172 | -35.0082 | -26.908
16.8233 | -35.0087 | -26.9156
16.8294 | -35.0092 | -26.9234
16.8356 | -35.0097 | -26.9312
16.8417 | -35.0102 | -26.9389
16.8478 | -35.0107 | -26.9464
16.8539 | -35.0112 | -26.954
16.86 | -35.0117 | -26.9619
16.8661 | -35.0122 | -26.9695
16.8722 | -35.0127 | -26.9772
16.8783 | -35.0132 | -26.9851
16.8844 | -35.0137 | -26.9924
16.8905 | -35.0142 | -27.0001
16.8966 | -35.0147 | -27.0078
16.9027 | -35.0153 | -27.0155
16.9088 | -35.0157 | -27.0232
16.9149 | -35.0163 | -27.031
16.921 | -35.0167 | -27.0385
16.9271 | -35.0172 | -27.0464
16.9332 | -35.0177 | -27.054
16.9393 | -35.0182 | -27.0618
16.9454 | -35.0187 | -27.0693
16.9515 | -35.0192 | -27.0772
16.9576 | -35.0197 | -27.0845
16.9637 | -35.0202 | -27.0922
16.9698 | -35.0207 | -27.0997
16.9759 | -35.0212 | -27.1073
16.9821 | -35.0217 | -27.115
16.9882 | -35.0222 | -27.1225
16.9943 | -35.0227 | -27.13
17.0004 | -35.0232 | -27.1376
17.0065 | -35.0238 | -27.1452
17.0126 | -35.0243 | -27.1529
17.0187 | -35.0248 | -27.1607
17.0248 | -35.0253 | -27.1683
17.0309 | -35.0258 | -27.1758
17.037 | -35.0263 | -27.1835
17.0431 | -35.0267 | -27.191
17.0492 | -35.0272 | -27.1987
17.0553 | -35.0278 | -27.2063
17.0614 | -35.0283 | -27.2138
17.0675 | -35.0288 | -27.2216
17.0736 | -35.0293 | -27.2293
17.0797 | -35.0298 | -27.2369
17.0858 | -35.0302 | -27.2447
17.0919 | -35.0308 | -27.2524
17.098 | -35.0312 | -27.2598
17.1041 | -35.0317 | -27.2672
17.1102 | -35.0322 | -27.2748
17.1163 | -35.0327 | -27.2824
17.1225 | -35.0333 | -27.2899
17.1286 | -35.0338 | -27.2974
17.1347 | -35.0343 | -27.3051
17.1408 | -35.0348 | -27.3126
17.1469 | -35.0353 | -27.3203
17.153 | -35.0358 | -27.3278
17.1591 | -35.0363 | -27.3354
17.1652 | -35.0368 | -27.3427
17.1713 | -35.0373 | -27.3501
17.1774 | -35.0378 | -27.3577
17.1835 | -35.0383 | -27.3652
17.1896 | -35.0388 | -27.3728
17.1957 | -35.0393 | -27.3804
17.2018 | -35.0398 | -27.3877
17.2079 | -35.0403 | -27.3953
17.214 | -35.0408 | -27.403
17.2201 | -35.0413 | -27.4106
17.2262 | -35.0418 | -27.4182
17.2323 | -35.0423 | -27.4259
17.2384 | -35.0428 | -27.4331
17.2445 | -35.0433 | -27.4406
17.2506 | -35.0438 | -27.4483
17.2567 | -35.0443 | -27.456
17.2628 | -35.0448 | -27.4635
17.269 | -35.0453 | -27.4709
17.2751 | -35.0458 | -27.4785
17.2812 | -35.0463 | -27.486
17.2873 | -35.0468 | -27.4936
17.2934 | -35.0473 | -27.5013
17.2995 | -35.0478 | -27.5088
17.3056 | -35.0483 | -27.5164
17.3117 | -35.0488 | -27.5241
17.3178 | -35.0493 | -27.5317
17.3239 | -35.0498 | -27.5391
17.33 | -35.0503 | -27.5466
17.3361 | -35.0508 | -27.5539
17.3422 | -35.0513 | -27.5612
17.3483 | -35.0518 | -27.5687
17.3544 | -35.0523 | -27.5763
17.3605 | -35.0528 | -27.5837
17.3666 | -35.0533 | -27.5912
17.3727 | -35.0538 | -27.5987
17.3788 | -35.0543 | -27.6059
17.3849 | -35.0548 | -27.6134
17.391 | -35.0553 | -27.621
17.3971 | -35.0558 | -27.6284
17.4032 | -35.0563 | -27.6361
17.4094 | -35.0568 | -27.6434
17.4155 | -35.0573 | -27.6509
17.4216 | -35.0578 | -27.6585
17.4277 | -35.0583 | -27.6661
17.4338 | -35.0588 | -27.6736
17.4399 | -35.0593 | -27.6811
17.446 | -35.0598 | -27.6885
17.4521 | -35.0603 | -27.696
17.4582 | -35.0608 | -27.7035
17.4643 | -35.0613 | -27.7109
17.4704 | -35.0618 | -27.7184
17.4765 | -35.0623 | -27.726
17.4826 | -35.0628 | -27.7334
17.4887 | -35.0633 | -27.7409
17.4948 | -35.0638 | -27.7485
17.5009 | -35.0643 | -27.756
17.507 | -35.0648 | -27.7635
17.5131 | -35.0653 | -27.771
17.5192 | -35.0658 | -27.7782
17.5253 | -35.0663 | -27.7857
17.5314 | -35.0667 | -27.7932
17.5375 | -35.0672 | -27.8007
17.5436 | -35.0677 | -27.8083
17.5497 | -35.0682 | -27.8158
17.5559 | -35.0687 | -27.8231
17.562 | -35.0692 | -27.8306
17.5681 | -35.0697 | -27.838
17.5742 | -35.0702 | -27.8455
17.5803 | -35.0707 | -27.8528
17.5864 | -35.0712 | -27.8601
17.5925 | -35.0717 | -27.8677
17.5986 | -35.0722 | -27.8752
17.6047 | -35.0727 | -27.8827
17.6108 | -35.0732 | -27.8901
17.6169 | -35.0737 | -27.8975
17.623 | -35.0742 | -27.9048
17.6291 | -35.0747 | -27.9122
17.6352 | -35.0752 | -27.9197
17.6413 | -35.0757 | -27.9271
17.6474 | -35.0762 | -27.9347
17.6535 | -35.0766 | -27.942
17.6596 | -35.0772 | -27.9493
17.6657 | -35.0777 | -27.9566
17.6718 | -35.0781 | -27.964
17.6779 | -35.0786 | -27.9714
17.684 | -35.0791 | -27.9788
17.6901 | -35.0796 | -27.9863
17.6963 | -35.0801 | -27.9934
17.7024 | -35.0806 | -28.0008
17.7085 | -35.0811 | -28.0081
17.7146 | -35.0816 | -28.0155
17.7207 | -35.0821 | -28.0229
17.7268 | -35.0826 | -28.0301
17.7329 | -35.0831 | -28.0376
17.739 | -35.0836 | -28.045
17.7451 | -35.0841 | -28.0524
17.7512 | -35.0846 | -28.0598
17.7573 | -35.0851 | -28.0673
17.7634 | -35.0856 | -28.0745
17.7695 | -35.0861 | -28.082
17.7756 | -35.0866 | -28.0894
17.7817 | -35.0871 | -28.0969
17.7878 | -35.0876 | -28.1046
17.7939 | -35.0881 | -28.112
17.8 | -35.0885 | -28.1192
17.8061 | -35.089 | -28.1266
17.8122 | -35.0895 | -28.1341
17.8183 | -35.09 | -28.1414
17.8244 | -35.0905 | -28.1489
17.8305 | -35.091 | -28.1563
17.8366 | -35.0915 | -28.1635
17.8428 | -35.092 | -28.171
17.8489 | -35.0925 | -28.1784
17.855 | -35.093 | -28.1856
17.8611 | -35.0935 | -28.1929
17.8672 | -35.094 | -28.2001
17.8733 | -35.0945 | -28.2076
17.8794 | -35.095 | -28.2149
17.8855 | -35.0955 | -28.2224
17.8916 | -35.096 | -28.2297
17.8977 | -35.0964 | -28.2373
17.9038 | -35.0969 | -28.2445
17.9099 | -35.0974 | -28.2517
17.916 | -35.0979 | -28.2591
17.9221 | -35.0984 | -28.2666
17.9282 | -35.0989 | -28.274
17.9343 | -35.0994 | -28.2814
17.9404 | -35.0999 | -28.2884
17.9465 | -35.1004 | -28.2957
17.9526 | -35.1009 | -28.3031
17.9587 | -35.1014 | -28.3104
17.9648 | -35.1019 | -28.3176
17.9709 | -35.1023 | -28.325
17.977 | -35.1028 | -28.3322
17.9832 | -35.1033 | -28.3395
17.9893 | -35.1038 | -28.3468
17.9954 | -35.1043 | -28.3542
18.0015 | -35.1048 | -28.3617
18.0076 | -35.1053 | -28.3689
18.0137 | -35.1058 | -28.3763
18.0198 | -35.1063 | -28.3835
18.0259 | -35.1068 | -28.3909
18.032 | -35.1073 | -28.3981
18.0381 | -35.1077 | -28.4056
18.0442 | -35.1082 | -28.4128
18.0503 | -35.1087 | -28.4201
18.0564 | -35.1092 | -28.4274
18.0625 | -35.1097 | -28.4349
18.0686 | -35.1102 | -28.4424
18.0747 | -35.1106 | -28.4498
18.0808 | -35.1111 | -28.4569
18.0869 | -35.1116 | -28.4643
18.093 | -35.1121 | -28.4715
18.0991 | -35.1126 | -28.4787
18.1052 | -35.1131 | -28.4861
18.1113 | -35.1136 | -28.4933
18.1174 | -35.1141 | -28.5004
18.1236 | -35.1146 | -28.5077
18.1297 | -35.115 | -28.5151
18.1358 | -35.1155 | -28.5224
18.1419 | -35.116 | -28.5297
18.148 | -35.1165 | -28.5371
18.1541 | -35.117 | -28.5443
18.1602 | -35.1175 | -28.5517
18.1663 | -35.118 | -28.5589
18.1724 | -35.1185 | -28.5663
18.1785 | -35.119 | -28.5735
18.1846 | -35.1194 | -28.5807
18.1907 | -35.1199 | -28.588
18.1968 | -35.1204 | -28.5954
18.2029 | -35.1209 | -28.6028
18.209 | -35.1214 | -28.6102
18.2151 | -35.1219 | -28.6176
18.2212 | -35.1223 | -28.6247
18.2273 | -35.1228 | -28.6321
18.2334 | -35.1233 | -28.6395
18.2395 | -35.1238 | -28.6468
18.2456 | -35.1243 | -28.6542
18.2517 | -35.1248 | -28.6615
18.2578 | -35.1253 | -28.6687
18.2639 | -35.1258 | -28.676
18.2701 | -35.1263 | -28.6833
18.2762 | -35.1268 | -28.6904
18.2823 | -35.1273 | -28.6977
18.2884 | -35.1278 | -28.7049
18.2945 | -35.1283 | -28.7119
18.3006 | -35.1288 | -28.7191
18.3067 | -35.1292 | -28.7265
18.3128 | -35.1297 | -28.7338
18.3189 | -35.1302 | -28.7412
18.325 | -35.1307 | -28.7482
18.3311 | -35.1312 | -28.7554
18.3372 | -35.1317 | -28.7628
18.3433 | -35.1321 | -28.77
18.3494 | -35.1326 | -28.7774
18.3555 | -35.1331 | -28.7846
18.3616 | -35.1336 | -28.7918
18.3677 | -35.1341 | -28.799
18.3738 | -35.1346 | -28.8061
18.3799 | -35.1351 | -28.8134
18.386 | -35.1356 | -28.8207
18.3921 | -35.1361 | -28.828
18.3982 | -35.1365 | -28.8351
18.4043 | -35.137 | -28.8423
18.4105 | -35.1375 | -28.8496
18.4166 | -35.1379 | -28.8567
18.4227 | -35.1384 | -28.864
18.4288 | -35.1389 | -28.8713
18.4349 | -35.1394 | -28.8786
18.441 | -35.1399 | -28.8858
18.4471 | -35.1403 | -28.8929
18.4532 | -35.1408 | -28.9002
18.4593 | -35.1413 | -28.9073
18.4654 | -35.1418 | -28.9145
18.4715 | -35.1423 | -28.9219
18.4776 | -35.1428 | -28.9291
18.4837 | -35.1432 | -28.9364
18.4898 | -35.1437 | -28.9436
18.4959 | -35.1442 | -28.9508
18.502 | -35.1446 | -28.9578
18.5081 | -35.1451 | -28.9651
18.5142 | -35.1456 | -28.9724
18.5203 | -35.1461 | -28.9796
18.5264 | -35.1465 | -28.9868
18.5325 | -35.147 | -28.994
18.5386 | -35.1475 | -29.0012
18.5447 | -35.148 | -29.0083
18.5508 | -35.1485 | -29.0154
18.557 | -35.149 | -29.0226
18.5631 | -35.1494 | -29.0296
18.5692 | -35.1499 | -29.0369
18.5753 | -35.1504 | -29.0436
18.5814 | -35.1509 | -29.0509
18.5875 | -35.1513 | -29.0579
18.5936 | -35.1518 | -29.0653
18.5997 | -35.1523 | -29.0725
18.6058 | -35.1528 | -29.0794
18.6119 | -35.1533 | -29.0866
18.618 | -35.1537 | -29.0937
18.6241 | -35.1542 | -29.1008
18.6302 | -35.1547 | -29.108
18.6363 | -35.1552 | -29.1153
18.6424 | -35.1557 | -29.1225
18.6485 | -35.1562 | -29.1295
18.6546 | -35.1566 | -29.1366
18.6607 | -35.1571 | -29.1437
18.6668 | -35.1576 | -29.1507
18.6729 | -35.158 | -29.1578
18.679 | -35.1585 | -29.1647
18.6851 | -35.159 | -29.1719
18.6912 | -35.1594 | -29.179
18.6974 | -35.1599 | -29.1861
18.7035 | -35.1604 | -29.1933
18.7096 | -35.1609 | -29.2006
18.7157 | -35.1614 | -29.2076
18.7218 | -35.1618 | -29.2148
18.7279 | -35.1623 | -29.222
18.734 | -35.1628 | -29.2293
18.7401 | -35.1633 | -29.2364
18.7462 | -35.1638 | -29.2434
18.7523 | -35.1643 | -29.2506
18.7584 | -35.1647 | -29.2579
18.7645 | -35.1652 | -29.2649
18.7706 | -35.1657 | -29.2721
18.7767 | -35.1662 | -29.2792
18.7828 | -35.1666 | -29.286
18.7889 | -35.1671 | -29.2932
18.795 | -35.1676 | -29.3004
18.8011 | -35.1681 | -29.3075
18.8072 | -35.1685 | -29.3148
18.8133 | -35.169 | -29.3219
18.8194 | -35.1695 | -29.3289
18.8255 | -35.17 | -29.3361
18.8316 | -35.1704 | -29.3433
18.8377 | -35.1709 | -29.3503
18.8439 | -35.1714 | -29.3574
18.85 | -35.1719 | -29.3645
18.8561 | -35.1724 | -29.3715
18.8622 | -35.1728 | -29.3788
18.8683 | -35.1733 | -29.386
18.8744 | -35.1738 | -29.393
18.8805 | -35.1743 | -29.4002
18.8866 | -35.1747 | -29.4074
18.8927 | -35.1752 | -29.4142
18.8988 | -35.1757 | -29.4213
18.9049 | -35.1762 | -29.4284
18.911 | -35.1767 | -29.4355
18.9171 | -35.1772 | -29.4426
18.9232 | -35.1776 | -29.4496
18.9293 | -35.1781 | -29.4567
18.9354 | -35.1786 | -29.4638
18.9415 | -35.1791 | -29.471
18.9476 | -35.1795 | -29.4782
18.9537 | -35.18 | -29.4853
18.9598 | -35.1805 | -29.492
18.9659 | -35.181 | -29.4992
18.972 | -35.1815 | -29.5063
18.9781 | -35.1819 | -29.5135
18.9843 | -35.1824 | -29.5207
18.9904 | -35.1829 | -29.5279
18.9965 | -35.1834 | -29.5347
19.0026 | -35.1838 | -29.5418
19.0087 | -35.1843 | -29.5489
19.0148 | -35.1848 | -29.5561
19.0209 | -35.1853 | -29.5633
19.027 | -35.1858 | -29.5706
19.0331 | -35.1862 | -29.5775
19.0392 | -35.1867 | -29.5846
19.0453 | -35.1872 | -29.5916
19.0514 | -35.1876 | -29.5987
19.0575 | -35.1881 | -29.606
19.0636 | -35.1886 | -29.6127
19.0697 | -35.1891 | -29.6197
19.0758 | -35.1895 | -29.6269
19.0819 | -35.19 | -29.634
19.088 | -35.1905 | -29.6412
19.0941 | -35.1909 | -29.6485
19.1002 | -35.1914 | -29.6553
19.1063 | -35.1919 | -29.6624
19.1124 | -35.1924 | -29.6695
19.1185 | -35.1928 | -29.6765
19.1246 | -35.1934 | -29.6837
19.1308 | -35.1938 | -29.6907
19.1369 | -35.1943 | -29.6977
19.143 | -35.1948 | -29.7048
19.1491 | -35.1952 | -29.7119
19.1552 | -35.1957 | -29.7191
19.1613 | -35.1962 | -29.7262
19.1674 | -35.1966 | -29.7332
19.1735 | -35.1971 | -29.7401
19.1796 | -35.1975 | -29.747
19.1857 | -35.198 | -29.7541
19.1918 | -35.1985 | -29.7612
19.1979 | -35.199 | -29.7684
19.204 | -35.1994 | -29.7755
19.2101 | -35.1999 | -29.7826
19.2162 | -35.2003 | -29.7896
19.2223 | -35.2008 | -29.7967
19.2284 | -35.2013 | -29.8038
19.2345 | -35.2017 | -29.8108
19.2406 | -35.2022 | -29.8177
19.2467 | -35.2027 | -29.8246
19.2528 | -35.2031 | -29.8317
19.2589 | -35.2036 | -29.8387
19.265 | -35.2041 | -29.846
19.2712 | -35.2045 | -29.853
19.2773 | -35.205 | -29.8601
19.2834 | -35.2055 | -29.8671
19.2895 | -35.206 | -29.8741
19.2956 | -35.2064 | -29.8812
19.3017 | -35.2069 | -29.8882
19.3078 | -35.2074 | -29.8954
19.3139 | -35.2079 | -29.9022
19.32 | -35.2083 | -29.9091
19.3261 | -35.2088 | -29.9162
19.3322 | -35.2093 | -29.923
19.3383 | -35.2098 | -29.93
19.3444 | -35.2103 | -29.9369
19.3505 | -35.2107 | -29.9439
19.3566 | -35.2112 | -29.9509
19.3627 | -35.2117 | -29.9581
19.3688 | -35.2121 | -29.9651
19.3749 | -35.2126 | -29.9721
19.381 | -35.2131 | -29.9789
19.3871 | -35.2136 | -29.9859
19.3932 | -35.214 | -29.993
19.3993 | -35.2145 | -30.0001
19.4054 | -35.215 | -30.0071
19.4115 | -35.2154 | -30.0144
19.4177 | -35.2159 | -30.0211
19.4238 | -35.2164 | -30.0281
19.4299 | -35.2168 | -30.035
19.436 | -35.2173 | -30.0421
19.4421 | -35.2178 | -30.049
19.4482 | -35.2182 | -30.056
19.4543 | -35.2187 | -30.0628
19.4604 | -35.2191 | -30.0699
19.4665 | -35.2196 | -30.0767
19.4726 | -35.2201 | -30.0837
19.4787 | -35.2205 | -30.0907
19.4848 | -35.221 | -30.0977
19.4909 | -35.2214 | -30.1044
19.497 | -35.2219 | -30.1114
19.5031 | -35.2224 | -30.1183
19.5092 | -35.2228 | -30.1253
19.5153 | -35.2233 | -30.1323
19.5214 | -35.2237 | -30.1389
19.5275 | -35.2242 | -30.1459
19.5336 | -35.2247 | -30.1529
19.5397 | -35.2252 | -30.16
19.5458 | -35.2256 | -30.167
19.5519 | -35.2261 | -30.174
19.5581 | -35.2265 | -30.181
19.5642 | -35.227 | -30.1879
19.5703 | -35.2274 | -30.195
19.5764 | -35.228 | -30.202
19.5825 | -35.2284 | -30.209
19.5886 | -35.2289 | -30.216
19.5947 | -35.2294 | -30.223
19.6008 | -35.2298 | -30.2298
19.6069 | -35.2303 | -30.2366
19.613 | -35.2308 | -30.2436
19.6191 | -35.2312 | -30.2505
19.6252 | -35.2317 | -30.2574
19.6313 | -35.2322 | -30.2639
19.6374 | -35.2326 | -30.271
19.6435 | -35.2331 | -30.2779
19.6496 | -35.2336 | -30.2847
19.6557 | -35.234 | -30.2917
19.6618 | -35.2345 | -30.2985
19.6679 | -35.235 | -30.3054
19.674 | -35.2354 | -30.3123
19.6801 | -35.2359 | -30.3193
19.6862 | -35.2364 | -30.3262
19.6923 | -35.2368 | -30.3331
19.6984 | -35.2373 | -30.3399
19.7046 | -35.2378 | -30.3469
19.7107 | -35.2382 | -30.3536
19.7168 | -35.2387 | -30.3606
19.7229 | -35.2392 | -30.3675
19.729 | -35.2396 | -30.3743
19.7351 | -35.2401 | -30.381
19.7412 | -35.2406 | -30.388
19.7473 | -35.2411 | -30.395
19.7534 | -35.2415 | -30.4017
19.7595 | -35.242 | -30.4086
19.7656 | -35.2424 | -30.4155
19.7717 | -35.2429 | -30.4222
19.7778 | -35.2434 | -30.4292
19.7839 | -35.2438 | -30.436
19.79 | -35.2443 | -30.443
19.7961 | -35.2448 | -30.45
19.8022 | -35.2452 | -30.4568
19.8083 | -35.2457 | -30.4638
19.8144 | -35.2462 | -30.4709
19.8205 | -35.2467 | -30.4777
19.8266 | -35.2471 | -30.4846
19.8327 | -35.2475 | -30.4915
19.8388 | -35.248 | -30.4982
19.845 | -35.2485 | -30.5052
19.8511 | -35.2489 | -30.5123
19.8572 | -35.2494 | -30.5192
19.8633 | -35.2498 | -30.5261
19.8694 | -35.2503 | -30.533
19.8755 | -35.2508 | -30.5394
19.8816 | -35.2512 | -30.5465
19.8877 | -35.2517 | -30.5534
19.8938 | -35.2521 | -30.5604
19.8999 | -35.2526 | -30.5671
19.906 | -35.253 | -30.5739
19.9121 | -35.2535 | -30.5806
19.9182 | -35.254 | -30.5875
19.9243 | -35.2544 | -30.5942
19.9304 | -35.2549 | -30.6012
19.9365 | -35.2554 | -30.6084
19.9426 | -35.2559 | -30.6149
19.9487 | -35.2563 | -30.6218
19.9548 | -35.2568 | -30.6286
19.9609 | -35.2572 | -30.6356
19.967 | -35.2577 | -30.6423
19.9731 | -35.2582 | -30.6493
19.9792 | -35.2586 | -30.6558
19.9853 | -35.2591 | -30.6626
19.9915 | -35.2596 | -30.6694
19.9976 | -35.26 | -30.6762
20.0037 | -35.2605 | -30.6831
20.0098 | -35.261 | -30.6899
20.0159 | -35.2614 | -30.6963
20.022 | -35.2619 | -30.7033
20.0281 | -35.2624 | -30.7101
20.0342 | -35.2628 | -30.7168
20.0403 | -35.2633 | -30.7236
20.0464 | -35.2638 | -30.7303
20.0525 | -35.2643 | -30.737
20.0586 | -35.2648 | -30.7439
20.0647 | -35.2653 | -30.7507
20.0708 | -35.2657 | -30.7576
20.0769 | -35.2662 | -30.7645
20.083 | -35.2666 | -30.7712
20.0891 | -35.2671 | -30.7782
20.0952 | -35.2676 | -30.7849
20.1013 | -35.268 | -30.7918
20.1074 | -35.2685 | -30.7987
20.1135 | -35.269 | -30.8058
20.1196 | -35.2694 | -30.8124
20.1257 | -35.2699 | -30.8193
20.1319 | -35.2703 | -30.8261
20.138 | -35.2708 | -30.833
20.1441 | -35.2713 | -30.8399
20.1502 | -35.2717 | -30.8469
20.1563 | -35.2722 | -30.8536
20.1624 | -35.2726 | -30.8603
20.1685 | -35.2731 | -30.8672
20.1746 | -35.2735 | -30.874
20.1807 | -35.274 | -30.8811
20.1868 | -35.2744 | -30.888
20.1929 | -35.2749 | -30.8947
20.199 | -35.2754 | -30.9015
20.2051 | -35.2758 | -30.9084
20.2112 | -35.2763 | -30.9153
20.2173 | -35.2768 | -30.9222
20.2234 | -35.2772 | -30.9292
20.2295 | -35.2777 | -30.9362
20.2356 | -35.2781 | -30.9434
20.2417 | -35.2786 | -30.9502
20.2478 | -35.279 | -30.9572
20.2539 | -35.2794 | -30.9641
20.26 | -35.2799 | -30.9708
20.2661 | -35.2804 | -30.9777
20.2723 | -35.2808 | -30.9846
20.2784 | -35.2813 | -30.9913
20.2845 | -35.2818 | -30.9981
20.2906 | -35.2823 | -31.0049
20.2967 | -35.2828 | -31.0115
20.3028 | -35.2832 | -31.0182
20.3089 | -35.2837 | -31.0249
20.315 | -35.2842 | -31.0316
20.3211 | -35.2846 | -31.0384
20.3272 | -35.2851 | -31.0453
20.3333 | -35.2856 | -31.0517
20.3394 | -35.286 | -31.0585
20.3455 | -35.2865 | -31.0653
20.3516 | -35.2869 | -31.0722
20.3577 | -35.2874 | -31.0787
20.3638 | -35.2878 | -31.0856
20.3699 | -35.2884 | -31.0922
20.376 | -35.2888 | -31.099
20.3821 | -35.2893 | -31.1057
20.3882 | -35.2897 | -31.1125
20.3943 | -35.2902 | -31.1191
20.4004 | -35.2906 | -31.1253
20.4065 | -35.2911 | -31.1319
20.4126 | -35.2915 | -31.1387
20.4188 | -35.292 | -31.1456
20.4249 | -35.2924 | -31.1525
20.431 | -35.2929 | -31.1593
20.4371 | -35.2933 | -31.1656
20.4432 | -35.2938 | -31.1722
20.4493 | -35.2942 | -31.1792
20.4554 | -35.2947 | -31.1859
20.4615 | -35.2952 | -31.1928
20.4676 | -35.2956 | -31.1997
20.4737 | -35.2961 | -31.2061
20.4798 | -35.2966 | -31.2129
20.4859 | -35.297 | -31.2198
20.492 | -35.2975 | -31.2266
20.4981 | -35.298 | -31.2333
20.5042 | -35.2984 | -31.24
20.5103 | -35.2989 | -31.2464
20.5164 | -35.2993 | -31.2532
20.5225 | -35.2998 | -31.2599
20.5286 | -35.3002 | -31.2666
20.5347 | -35.3007 | -31.2733
20.5408 | -35.3012 | -31.2796
20.5469 | -35.3016 | -31.2864
20.553 | -35.3021 | -31.2933
20.5592 | -35.3026 | -31.2999
20.5653 | -35.303 | -31.3067
20.5714 | -35.3035 | -31.3135
20.5775 | -35.304 | -31.3201
20.5836 | -35.3044 | -31.327
20.5897 | -35.3049 | -31.3336
20.5958 | -35.3054 | -31.3405
20.6019 | -35.3058 | -31.3472
20.608 | -35.3063 | -31.354
20.6141 | -35.3068 | -31.3606
20.6202 | -35.3072 | -31.3675
20.6263 | -35.3077 | -31.374
20.6324 | -35.3082 | -31.3809
20.6385 | -35.3086 | -31.3874
20.6446 | -35.3091 | -31.3939
20.6507 | -35.3095 | -31.4005
20.6568 | -35.31 | -31.4072
20.6629 | -35.3104 | -31.4139
20.669 | -35.3109 | -31.4207
20.6751 | -35.3113 | -31.4273
20.6812 | -35.3118 | -31.4338
20.6873 | -35.3122 | -31.4407
20.6934 | -35.3127 | -31.4474
20.6995 | -35.3132 | -31.454
20.7057 | -35.3136 | -31.4607
20.7118 | -35.3141 | -31.4672
20.7179 | -35.3146 | -31.4736
20.724 | -35.3151 | -31.4803
20.7301 | -35.3155 | -31.4869
20.7362 | -35.316 | -31.4936
20.7423 | -35.3165 | -31.5003
20.7484 | -35.3169 | -31.5069
20.7545 | -35.3174 | -31.5135
20.7606 | -35.3179 | -31.5201
20.7667 | -35.3184 | -31.5268
20.7728 | -35.3189 | -31.5333
20.7789 | -35.3193 | -31.5404
20.785 | -35.3198 | -31.5468
20.7911 | -35.3203 | -31.5531
20.7972 | -35.3207 | -31.5599
20.8033 | -35.3212 | -31.5667
20.8094 | -35.3217 | -31.5732
20.8155 | -35.3221 | -31.5799
20.8216 | -35.3226 | -31.5865
20.8277 | -35.323 | -31.593
20.8338 | -35.3235 | -31.5995
20.8399 | -35.324 | -31.6063
20.8461 | -35.3244 | -31.6128
20.8522 | -35.3249 | -31.6195
20.8583 | -35.3253 | -31.6259
20.8644 | -35.3258 | -31.6326
20.8705 | -35.3263 | -31.6393
20.8766 | -35.3267 | -31.6461
20.8827 | -35.3272 | -31.6527
20.8888 | -35.3276 | -31.6595
20.8949 | -35.3281 | -31.6657
20.901 | -35.3285 | -31.6724
20.9071 | -35.329 | -31.6791
20.9132 | -35.3295 | -31.6857
20.9193 | -35.3299 | -31.6923
20.9254 | -35.3304 | -31.6989
20.9315 | -35.3309 | -31.7055
20.9376 | -35.3314 | -31.7123
20.9437 | -35.3318 | -31.7191
20.9498 | -35.3323 | -31.7258
20.9559 | -35.3328 | -31.7324
20.962 | -35.3332 | -31.7391
20.9681 | -35.3337 | -31.7456
20.9742 | -35.3342 | -31.7522
20.9803 | -35.3346 | -31.759
20.9864 | -35.3351 | -31.7658
20.9926 | -35.3355 | -31.7725
20.9987 | -35.336 | -31.7791
21.0048 | -35.3365 | -31.7857
21.0109 | -35.337 | -31.7924
21.017 | -35.3375 | -31.799
21.0231 | -35.338 | -31.8057
21.0292 | -35.3384 | -31.8123
21.0353 | -35.3389 | -31.8189
21.0414 | -35.3394 | -31.8255
21.0475 | -35.3398 | -31.8325
21.0536 | -35.3403 | -31.8392
21.0597 | -35.3407 | -31.8457
21.0658 | -35.3412 | -31.8522
21.0719 | -35.3417 | -31.8588
21.078 | -35.3421 | -31.8655
21.0841 | -35.3426 | -31.8722
21.0902 | -35.3431 | -31.8791
21.0963 | -35.3436 | -31.8861
21.1024 | -35.344 | -31.8928
21.1085 | -35.3445 | -31.8995
21.1146 | -35.345 | -31.9063
21.1207 | -35.3455 | -31.913
21.1268 | -35.346 | -31.9195
21.133 | -35.3464 | -31.9263
21.1391 | -35.3469 | -31.9327
21.1452 | -35.3474 | -31.9394
21.1513 | -35.3478 | -31.9459
21.1574 | -35.3483 | -31.9524
21.1635 | -35.3488 | -31.959
21.1696 | -35.3492 | -31.9655
21.1757 | -35.3497 | -31.9718
21.1818 | -35.3502 | -31.9787
21.1879 | -35.3506 | -31.9855
21.194 | -35.3511 | -31.9925
21.2001 | -35.3516 | -31.9988
21.2062 | -35.3521 | -32.0052
21.2123 | -35.3526 | -32.0116
21.2184 | -35.353 | -32.018
21.2245 | -35.3535 | -32.0254
21.2306 | -35.3541 | -32.033
21.2367 | -35.3546 | -32.0403
21.2428 | -35.3552 | -32.0469
21.2489 | -35.3557 | -32.0531
21.255 | -35.3562 | -32.0597
21.2611 | -35.3566 | -32.0665
21.2672 | -35.3571 | -32.0729
21.2733 | -35.3576 | -32.0799
21.2795 | -35.358 | -32.0863
21.2856 | -35.3585 | -32.093
21.2917 | -35.359 | -32.0994
21.2978 | -35.3594 | -32.106
21.3039 | -35.3599 | -32.1127
21.31 | -35.3604 | -32.1195
21.3161 | -35.3608 | -32.1261
21.3222 | -35.3613 | -32.133
21.3283 | -35.3617 | -32.1394
21.3344 | -35.3622 | -32.1458
21.3405 | -35.3627 | -32.1523
21.3466 | -35.3632 | -32.159
21.3527 | -35.3638 | -32.1649
21.3588 | -35.3643 | -32.1717
21.3649 | -35.3647 | -32.1781
21.371 | -35.3652 | -32.1846
21.3771 | -35.3657 | -32.191
21.3832 | -35.3661 | -32.1977
21.3893 | -35.3666 | -32.2042
21.3954 | -35.367 | -32.2107
21.4015 | -35.3675 | -32.2172
21.4076 | -35.368 | -32.2238
21.4137 | -35.3684 | -32.2305
21.4199 | -35.3689 | -32.2371
21.426 | -35.3694 | -32.2435
21.4321 | -35.3698 | -32.2501
21.4382 | -35.3703 | -32.2564
21.4443 | -35.3708 | -32.263
21.4504 | -35.3713 | -32.2698
21.4565 | -35.3718 | -32.2761
21.4626 | -35.3722 | -32.2827
21.4687 | -35.3727 | -32.2892
21.4748 | -35.3732 | -32.2958
21.4809 | -35.3737 | -32.3026
21.487 | -35.3741 | -32.3092
21.4931 | -35.3746 | -32.3157
21.4992 | -35.3751 | -32.3223
21.5053 | -35.3755 | -32.329
21.5114 | -35.376 | -32.3356
21.5175 | -35.3764 | -32.3423
21.5236 | -35.3769 | -32.3487
21.5297 | -35.3774 | -32.3555
21.5358 | -35.3778 | -32.3619
21.5419 | -35.3783 | -32.3687
21.548 | -35.3787 | -32.3754
21.5541 | -35.3792 | -32.382
21.5602 | -35.3797 | -32.3887
21.5664 | -35.3802 | -32.3946
21.5725 | -35.3807 | -32.4012
21.5786 | -35.3811 | -32.4079
21.5847 | -35.3816 | -32.4143
21.5908 | -35.3821 | -32.4208
21.5969 | -35.3826 | -32.4276
21.603 | -35.383 | -32.4342
21.6091 | -35.3835 | -32.4409
21.6152 | -35.384 | -32.4475
21.6213 | -35.3844 | -32.4541
21.6274 | -35.3849 | -32.4605
21.6335 | -35.3854 | -32.4669
21.6396 | -35.3859 | -32.4736
21.6457 | -35.3863 | -32.4805
21.6518 | -35.3868 | -32.4871
21.6579 | -35.3873 | -32.4935
21.664 | -35.3878 | -32.5001
21.6701 | -35.3883 | -32.5063
21.6762 | -35.3888 | -32.5128
21.6823 | -35.3892 | -32.5192
21.6884 | -35.3897 | -32.5257
21.6945 | -35.3902 | -32.5319
21.7006 | -35.3907 | -32.5382
21.7068 | -35.3912 | -32.5444
21.7129 | -35.3916 | -32.5507
21.719 | -35.3921 | -32.5574
21.7251 | -35.3926 | -32.564
21.7312 | -35.393 | -32.5706
21.7373 | -35.3935 | -32.5769
21.7434 | -35.394 | -32.5833
21.7495 | -35.3945 | -32.59
21.7556 | -35.3949 | -32.5967
21.7617 | -35.3954 | -32.6033
21.7678 | -35.3959 | -32.6098
21.7739 | -35.3963 | -32.6161
21.78 | -35.3969 | -32.6226
21.7861 | -35.3973 | -32.6292
21.7922 | -35.3978 | -32.6357
21.7983 | -35.3983 | -32.6422
21.8044 | -35.3988 | -32.6488
21.8105 | -35.3993 | -32.6554
21.8166 | -35.3998 | -32.6618
21.8227 | -35.4002 | -32.6681
21.8288 | -35.4007 | -32.6747
21.8349 | -35.4012 | -32.6813
21.841 | -35.4016 | -32.6879
21.8471 | -35.4021 | -32.6941
21.8533 | -35.4026 | -32.7006
21.8594 | -35.4031 | -32.7072
21.8655 | -35.4035 | -32.7138
21.8716 | -35.404 | -32.7205
21.8777 | -35.4045 | -32.727
21.8838 | -35.405 | -32.7336
21.8899 | -35.4055 | -32.7403
21.896 | -35.406 | -32.7469
21.9021 | -35.4064 | -32.7533
21.9082 | -35.4069 | -32.76
21.9143 | -35.4074 | -32.7659
21.9204 | -35.4079 | -32.7725
21.9265 | -35.4084 | -32.7788
21.9326 | -35.4088 | -32.785
21.9387 | -35.4093 | -32.7912
21.9448 | -35.4098 | -32.798
21.9509 | -35.4102 | -32.8044
21.957 | -35.4107 | -32.8111
21.9631 | -35.4112 | -32.8175
21.9692 | -35.4116 | -32.8241
21.9753 | -35.4121 | -32.8307
21.9814 | -35.4126 | -32.837
21.9875 | -35.4131 | -32.8433
21.9937 | -35.4136 | -32.8499
21.9998 | -35.4141 | -32.8564
22.0059 | -35.4145 | -32.863
22.012 | -35.4151 | -32.8695
22.0181 | -35.4156 | -32.8762
22.0242 | -35.416 | -32.8826
22.0303 | -35.4165 | -32.8893
22.0364 | -35.417 | -32.8957
22.0425 | -35.4175 | -32.9023
22.0486 | -35.4179 | -32.9086
22.0547 | -35.4184 | -32.9145
22.0608 | -35.4188 | -32.9209
22.0669 | -35.4193 | -32.9273
22.073 | -35.4198 | -32.9338
22.0791 | -35.4202 | -32.9406
22.0852 | -35.4207 | -32.9469
22.0913 | -35.4212 | -32.9533
22.0974 | -35.4216 | -32.9597
22.1035 | -35.4221 | -32.9664
22.1096 | -35.4226 | -32.9729
22.1157 | -35.423 | -32.9792
22.1218 | -35.4235 | -32.9857
22.1279 | -35.4239 | -32.992
22.134 | -35.4244 | -32.9987
22.1402 | -35.4249 | -33.005
22.1463 | -35.4254 | -33.0114
22.1524 | -35.4258 | -33.0178
22.1585 | -35.4263 | -33.0244
22.1646 | -35.4268 | -33.0306
22.1707 | -35.4272 | -33.0371
22.1768 | -35.4277 | -33.0435
22.1829 | -35.4282 | -33.0499
22.189 | -35.4287 | -33.0564
22.1951 | -35.4292 | -33.0625
22.2012 | -35.4297 | -33.0689
22.2073 | -35.4302 | -33.0754
22.2134 | -35.4306 | -33.0818
22.2195 | -35.4311 | -33.0882
22.2256 | -35.4316 | -33.0948
22.2317 | -35.4321 | -33.1009
22.2378 | -35.4326 | -33.1073
22.2439 | -35.4331 | -33.1137
22.25 | -35.4335 | -33.12
22.2561 | -35.4339 | -33.1265
22.2622 | -35.4344 | -33.133
22.2683 | -35.4349 | -33.1393
22.2744 | -35.4354 | -33.1459
22.2806 | -35.4359 | -33.1523
22.2867 | -35.4363 | -33.1586
22.2928 | -35.4368 | -33.1651
22.2989 | -35.4373 | -33.1716
22.305 | -35.4378 | -33.1777
22.3111 | -35.4383 | -33.1843
22.3172 | -35.4388 | -33.1907
22.3233 | -35.4392 | -33.1972
22.3294 | -35.4397 | -33.2036
22.3355 | -35.4402 | -33.21
22.3416 | -35.4407 | -33.2165
22.3477 | -35.4412 | -33.2228
22.3538 | -35.4416 | -33.2292
22.3599 | -35.4421 | -33.2356
22.366 | -35.4425 | -33.2419
22.3721 | -35.443 | -33.2481
22.3782 | -35.4435 | -33.2546
22.3843 | -35.4439 | -33.2611
22.3904 | -35.4444 | -33.2674
22.3965 | -35.4449 | -33.2737
22.4026 | -35.4454 | -33.28
22.4087 | -35.4458 | -33.2864
22.4148 | -35.4463 | -33.2929
22.4209 | -35.4469 | -33.2995
22.4271 | -35.4473 | -33.3057
22.4332 | -35.4478 | -33.3123
22.4393 | -35.4483 | -33.3186
22.4454 | -35.4488 | -33.325
22.4515 | -35.4492 | -33.3313
22.4576 | -35.4497 | -33.3376
22.4637 | -35.4501 | -33.3441
22.4698 | -35.4505 | -33.3505
22.4759 | -35.451 | -33.3563
22.482 | -35.4515 | -33.3631
22.4881 | -35.4519 | -33.3694
22.4942 | -35.4524 | -33.3758
22.5003 | -35.4529 | -33.3822
22.5064 | -35.4533 | -33.3885
22.5125 | -35.4538 | -33.3946
22.5186 | -35.4543 | -33.4009
22.5247 | -35.4548 | -33.4071
22.5308 | -35.4552 | -33.4135
22.5369 | -35.4557 | -33.4199
22.543 | -35.4561 | -33.4265
22.5491 | -35.4566 | -33.4326
22.5552 | -35.4571 | -33.439
22.5613 | -35.4576 | -33.4456
22.5675 | -35.4581 | -33.4519
22.5736 | -35.4585 | -33.4583
22.5797 | -35.459 | -33.4646
22.5858 | -35.4595 | -33.4707
22.5919 | -35.46 | -33.477
22.598 | -35.4604 | -33.4834
22.6041 | -35.4609 | -33.4898
22.6102 | -35.4614 | -33.4961
22.6163 | -35.4619 | -33.5022
22.6224 | -35.4623 | -33.5085
22.6285 | -35.4628 | -33.5146
22.6346 | -35.4633 | -33.5211
22.6407 | -35.4638 | -33.5275
22.6468 | -35.4643 | -33.5339
22.6529 | -35.4647 | -33.5401
22.659 | -35.4652 | -33.5463
22.6651 | -35.4657 | -33.5527
22.6712 | -35.4661 | -33.559
22.6773 | -35.4666 | -33.5652
22.6834 | -35.4671 | -33.5716
22.6895 | -35.4676 | -33.578
22.6956 | -35.468 | -33.5843
22.7017 | -35.4685 | -33.5905
22.7079 | -35.469 | -33.5966
22.714 | -35.4695 | -33.6031
22.7201 | -35.4699 | -33.6095
22.7262 | -35.4703 | -33.6156
22.7323 | -35.4708 | -33.6217
22.7384 | -35.4713 | -33.6281
22.7445 | -35.4718 | -33.6343
22.7506 | -35.4722 | -33.6404
22.7567 | -35.4727 | -33.6468
22.7628 | -35.4732 | -33.6531
22.7689 | -35.4736 | -33.6595
22.775 | -35.4741 | -33.666
22.7811 | -35.4746 | -33.6724
22.7872 | -35.4751 | -33.6788
22.7933 | -35.4755 | -33.6852
22.7994 | -35.476 | -33.6914
22.8055 | -35.4766 | -33.6977
22.8116 | -35.477 | -33.7039
22.8177 | -35.4775 | -33.7105
22.8238 | -35.478 | -33.7167
22.8299 | -35.4785 | -33.7229
22.836 | -35.4789 | -33.7294
22.8421 | -35.4794 | -33.7358
22.8482 | -35.4798 | -33.7422
22.8544 | -35.4803 | -33.7487
22.8605 | -35.4807 | -33.7549
22.8666 | -35.4812 | -33.7608
22.8727 | -35.4816 | -33.7668
22.8788 | -35.4821 | -33.7733
22.8849 | -35.4826 | -33.7796
22.891 | -35.483 | -33.7859
22.8971 | -35.4835 | -33.792
22.9032 | -35.484 | -33.7981
22.9093 | -35.4844 | -33.8045
22.9154 | -35.4848 | -33.811
22.9215 | -35.4853 | -33.8173
22.9276 | -35.4857 | -33.8236
22.9337 | -35.4862 | -33.8297
22.9398 | -35.4867 | -33.8359
22.9459 | -35.4871 | -33.8424
22.952 | -35.4876 | -33.8486
22.9581 | -35.488 | -33.855
22.9642 | -35.4885 | -33.8611
22.9703 | -35.4889 | -33.8671
22.9764 | -35.4894 | -33.8732
22.9825 | -35.4898 | -33.8793
22.9886 | -35.4903 | -33.8859
22.9948 | -35.4908 | -33.8924
23.0009 | -35.4912 | -33.8987
23.007 | -35.4917 | -33.9048
23.0131 | -35.4922 | -33.9112
23.0192 | -35.4926 | -33.9174
23.0253 | -35.4931 | -33.9236
23.0314 | -35.4935 | -33.9301
23.0375 | -35.494 | -33.9366
23.0436 | -35.4944 | -33.9424
23.0497 | -35.4948 | -33.9486
23.0558 | -35.4953 | -33.955
23.0619 | -35.4957 | -33.9614
23.068 | -35.4962 | -33.9677
23.0741 | -35.4966 | -33.9736
23.0802 | -35.4971 | -33.9799
23.0863 | -35.4975 | -33.9861
23.0924 | -35.498 | -33.9923
23.0985 | -35.4984 | -33.9986
23.1046 | -35.4989 | -34.0047
23.1107 | -35.4993 | -34.0104
23.1168 | -35.4998 | -34.0168
23.1229 | -35.5002 | -34.0231
23.129 | -35.5007 | -34.0292
23.1351 | -35.5011 | -34.0354
23.1413 | -35.5016 | -34.0417
23.1474 | -35.5021 | -34.0478
23.1535 | -35.5025 | -34.0541
23.1596 | -35.503 | -34.0603
23.1657 | -35.5034 | -34.0665
23.1718 | -35.5039 | -34.0727
23.1779 | -35.5044 | -34.079
23.184 | -35.5048 | -34.0849
23.1901 | -35.5053 | -34.0911
23.1962 | -35.5058 | -34.0975
23.2023 | -35.5063 | -34.1039
23.2084 | -35.5067 | -34.1098
23.2145 | -35.5072 | -34.1158
23.2206 | -35.5076 | -34.1223
23.2267 | -35.5082 | -34.1287
23.2328 | -35.5086 | -34.135
23.2389 | -35.5091 | -34.1413
23.245 | -35.5095 | -34.1474
23.2511 | -35.51 | -34.1532
23.2572 | -35.5105 | -34.1597
23.2633 | -35.5109 | -34.1659
23.2694 | -35.5113 | -34.172
23.2755 | -35.5118 | -34.1783
23.2817 | -35.5123 | -34.1845
23.2878 | -35.5127 | -34.1905
23.2939 | -35.5132 | -34.1966
23.3 | -35.5136 | -34.203
23.3061 | -35.5141 | -34.2093
23.3122 | -35.5146 | -34.2155
23.3183 | -35.515 | -34.2217
23.3244 | -35.5154 | -34.2275
23.3305 | -35.5159 | -34.2337
23.3366 | -35.5163 | -34.2401
23.3427 | -35.5168 | -34.2462
23.3488 | -35.5172 | -34.2523
23.3549 | -35.5177 | -34.2586
23.361 | -35.5181 | -34.2645
23.3671 | -35.5186 | -34.2706
23.3732 | -35.5191 | -34.2768
23.3793 | -35.5195 | -34.283
23.3854 | -35.52 | -34.2893
23.3915 | -35.5205 | -34.2952
23.3976 | -35.5209 | -34.3014
23.4037 | -35.5214 | -34.3075
23.4098 | -35.5218 | -34.3138
23.4159 | -35.5223 | -34.3199
23.422 | -35.5227 | -34.3261
23.4282 | -35.5232 | -34.3318
23.4343 | -35.5237 | -34.338
23.4404 | -35.5241 | -34.3443
23.4465 | -35.5245 | -34.3505
23.4526 | -35.525 | -34.3567
23.4587 | -35.5254 | -34.3628
23.4648 | -35.5259 | -34.3686
23.4709 | -35.5263 | -34.3749
23.477 | -35.5267 | -34.3811
23.4831 | -35.5272 | -34.3872
23.4892 | -35.5276 | -34.3933
23.4953 | -35.5281 | -34.3995
23.5014 | -35.5285 | -34.4053
23.5075 | -35.5289 | -34.4115
23.5136 | -35.5294 | -34.4176
23.5197 | -35.5298 | -34.4235
23.5258 | -35.5303 | -34.4297
23.5319 | -35.5307 | -34.4355
23.538 | -35.5312 | -34.4416
23.5441 | -35.5316 | -34.4478
23.5502 | -35.5321 | -34.4539
23.5563 | -35.5325 | -34.46
23.5624 | -35.5329 | -34.4661
23.5686 | -35.5334 | -34.472
23.5747 | -35.5338 | -34.4782
23.5808 | -35.5342 | -34.4841
23.5869 | -35.5347 | -34.4903
23.593 | -35.5351 | -34.4965
23.5991 | -35.5355 | -34.5026
23.6052 | -35.536 | -34.5086
23.6113 | -35.5364 | -34.5148
23.6174 | -35.5369 | -34.521
23.6235 | -35.5373 | -34.527
23.6296 | -35.5378 | -34.5332
23.6357 | -35.5382 | -34.5392
23.6418 | -35.5387 | -34.545
23.6479 | -35.5391 | -34.5512
23.654 | -35.5396 | -34.5573
23.6601 | -35.54 | -34.5634
23.6662 | -35.5404 | -34.5693
23.6723 | -35.5408 | -34.5754
23.6784 | -35.5413 | -34.5814
23.6845 | -35.5417 | -34.5876
23.6906 | -35.5422 | -34.5938
23.6967 | -35.5426 | -34.6
23.7028 | -35.543 | -34.6061
23.7089 | -35.5434 | -34.6118
23.7151 | -35.5438 | -34.6178
23.7212 | -35.5443 | -34.6238
23.7273 | -35.5447 | -34.63
23.7334 | -35.5452 | -34.6364
23.7395 | -35.5456 | -34.6424
23.7456 | -35.546 | -34.6482
23.7517 | -35.5464 | -34.6541
23.7578 | -35.5469 | -34.6602
23.7639 | -35.5473 | -34.6662
23.77 | -35.5477 | -34.6723
23.7761 | -35.5482 | -34.6783
23.7822 | -35.5486 | -34.6842
23.7883 | -35.5491 | -34.6901
23.7944 | -35.5495 | -34.6961
23.8005 | -35.5499 | -34.7021
23.8066 | -35.5504 | -34.7083
23.8127 | -35.5507 | -34.7144
23.8188 | -35.5512 | -34.7207
23.8249 | -35.5517 | -34.7268
23.831 | -35.5521 | -34.7328
23.8371 | -35.5526 | -34.7386
23.8432 | -35.553 | -34.7448
23.8493 | -35.5534 | -34.7504
23.8555 | -35.5538 | -34.7565
23.8616 | -35.5543 | -34.7626
23.8677 | -35.5547 | -34.7687
23.8738 | -35.5551 | -34.7745
23.8799 | -35.5556 | -34.7806
23.886 | -35.556 | -34.7867
23.8921 | -35.5564 | -34.7926
23.8982 | -35.5568 | -34.7986
23.9043 | -35.5573 | -34.805
23.9104 | -35.5577 | -34.8111
23.9165 | -35.5581 | -34.8172
23.9226 | -35.5585 | -34.8229
23.9287 | -35.559 | -34.8288
23.9348 | -35.5594 | -34.8349
23.9409 | -35.5598 | -34.8409
23.947 | -35.5602 | -34.8469
23.9531 | -35.5606 | -34.8526
23.9592 | -35.5611 | -34.8587
23.9653 | -35.5615 | -34.865
23.9714 | -35.5619 | -34.8712
23.9775 | -35.5623 | -34.8773
23.9836 | -35.5628 | -34.8832
23.9897 | -35.5632 | -34.8891
23.9958 | -35.5636 | -34.8952
24.002 | -35.564 | -34.9011
24.0081 | -35.5645 | -34.9072
24.0142 | -35.5649 | -34.9131
24.0203 | -35.5653 | -34.919
24.0264 | -35.5658 | -34.9247
24.0325 | -35.5662 | -34.9307
24.0386 | -35.5666 | -34.9369
24.0447 | -35.567 | -34.9429
24.0508 | -35.5674 | -34.9488
24.0569 | -35.5679 | -34.9549
24.063 | -35.5683 | -34.9605
24.0691 | -35.5687 | -34.9667
24.0752 | -35.5691 | -34.9726
24.0813 | -35.5696 | -34.9783
24.0874 | -35.57 | -34.9843
24.0935 | -35.5704 | -34.9903
24.0996 | -35.5708 | -34.9958
24.1057 | -35.5712 | -35.0019
24.1118 | -35.5716 | -35.008
24.1179 | -35.5721 | -35.0139
24.124 | -35.5725 | -35.02
24.1301 | -35.5729 | -35.0259
24.1362 | -35.5733 | -35.032
24.1424 | -35.5738 | -35.038
24.1485 | -35.5742 | -35.0442
24.1546 | -35.5746 | -35.0501
24.1607 | -35.575 | -35.0563
24.1668 | -35.5754 | -35.0619
24.1729 | -35.5758 | -35.0679
24.179 | -35.5762 | -35.0738
24.1851 | -35.5766 | -35.0795
24.1912 | -35.5771 | -35.0855
24.1973 | -35.5775 | -35.0914
24.2034 | -35.5779 | -35.0968
24.2095 | -35.5783 | -35.1028
24.2156 | -35.5787 | -35.109
24.2217 | -35.5792 | -35.115
24.2278 | -35.5796 | -35.1209
24.2339 | -35.58 | -35.1266
24.24 | -35.5804 | -35.1322
24.2461 | -35.5809 | -35.138
24.2522 | -35.5813 | -35.1439
24.2583 | -35.5817 | -35.1499
24.2644 | -35.5821 | -35.1558
24.2705 | -35.5825 | -35.1615
24.2766 | -35.583 | -35.1674
24.2827 | -35.5834 | -35.1733
24.2889 | -35.5838 | -35.1793
24.295 | -35.5842 | -35.1851
24.3011 | -35.5846 | -35.1912
24.3072 | -35.585 | -35.197
24.3133 | -35.5854 | -35.2029
24.3194 | -35.5858 | -35.2089
24.3255 | -35.5862 | -35.2149
24.3316 | -35.5866 | -35.2207
24.3377 | -35.5871 | -35.2265
24.3438 | -35.5874 | -35.232
24.3499 | -35.5878 | -35.2379
24.356 | -35.5882 | -35.2439
24.3621 | -35.5886 | -35.2499
24.3682 | -35.589 | -35.2556
24.3743 | -35.5894 | -35.2617
24.3804 | -35.5898 | -35.2674
24.3865 | -35.5902 | -35.273
24.3926 | -35.5906 | -35.2791
24.3987 | -35.591 | -35.285
24.4048 | -35.5914 | -35.2907
24.4109 | -35.5918 | -35.2965
24.417 | -35.5922 | -35.3024
24.4231 | -35.5926 | -35.3082
24.4293 | -35.593 | -35.314
24.4354 | -35.5935 | -35.3199
24.4415 | -35.5939 | -35.3257
24.4476 | -35.5943 | -35.3313
24.4537 | -35.5947 | -35.337
24.4598 | -35.5951 | -35.3428
24.4659 | -35.5955 | -35.3487
24.472 | -35.5959 | -35.3546
24.4781 | -35.5963 | -35.3605
24.4842 | -35.5966 | -35.3661
24.4903 | -35.5971 | -35.3722
24.4964 | -35.5975 | -35.378
24.5025 | -35.5978 | -35.384
24.5086 | -35.5983 | -35.3895
24.5147 | -35.5987 | -35.3953
24.5208 | -35.5991 | -35.4012
24.5269 | -35.5995 | -35.4071
24.533 | -35.5999 | -35.413
24.5391 | -35.6003 | -35.4188
24.5452 | -35.6007 | -35.4245
24.5513 | -35.6011 | -35.43
24.5574 | -35.6015 | -35.4359
24.5635 | -35.6019 | -35.4419
24.5696 | -35.6023 | -35.4477
24.5758 | -35.6027 | -35.4536
24.5819 | -35.6031 | -35.4595
24.588 | -35.6035 | -35.4652
24.5941 | -35.6039 | -35.471
24.6002 | -35.6043 | -35.4767
24.6063 | -35.6047 | -35.4825
24.6124 | -35.6051 | -35.4882
24.6185 | -35.6055 | -35.494
24.6246 | -35.6059 | -35.4994
24.6307 | -35.6063 | -35.5053
24.6368 | -35.6067 | -35.5112
24.6429 | -35.607 | -35.5171
24.649 | -35.6075 | -35.5229
24.6551 | -35.6078 | -35.5286
24.6612 | -35.6082 | -35.5342
24.6673 | -35.6086 | -35.5402
24.6734 | -35.609 | -35.5461
24.6795 | -35.6094 | -35.552
24.6856 | -35.6098 | -35.5579
24.6917 | -35.6102 | -35.5637
24.6978 | -35.6106 | -35.5691
24.7039 | -35.611 | -35.575
24.71 | -35.6114 | -35.5809
24.7162 | -35.6118 | -35.5868
24.7223 | -35.6122 | -35.5924
24.7284 | -35.6125 | -35.5981
24.7345 | -35.6129 | -35.6038
24.7406 | -35.6133 | -35.6095
24.7467 | -35.6137 | -35.6154
24.7528 | -35.6141 | -35.6211
24.7589 | -35.6145 | -35.6269
24.765 | -35.6149 | -35.6324
24.7711 | -35.6152 | -35.6382
24.7772 | -35.6156 | -35.6439
24.7833 | -35.616 | -35.6496
24.7894 | -35.6164 | -35.6555
24.7955 | -35.6168 | -35.6613
24.8016 | -35.6172 | -35.6669
24.8077 | -35.6175 | -35.6726
24.8138 | -35.6179 | -35.6782
24.8199 | -35.6183 | -35.6841
24.826 | -35.6187 | -35.69
24.8321 | -35.6191 | -35.6957
24.8382 | -35.6195 | -35.7014
24.8443 | -35.6199 | -35.7072
24.8504 | -35.6203 | -35.713
24.8565 | -35.6206 | -35.7187
24.8627 | -35.621 | -35.7245
24.8688 | -35.6213 | -35.7303
24.8749 | -35.6217 | -35.7359
24.881 | -35.6221 | -35.7417
24.8871 | -35.6225 | -35.7474
24.8932 | -35.6229 | -35.7532
24.8993 | -35.6233 | -35.7591
24.9054 | -35.6237 | -35.7647
24.9115 | -35.6241 | -35.7704
24.9176 | -35.6245 | -35.7761
24.9237 | -35.6248 | -35.782
24.9298 | -35.6252 | -35.7879
24.9359 | -35.6256 | -35.7938
24.942 | -35.626 | -35.7995
24.9481 | -35.6264 | -35.8052
24.9542 | -35.6267 | -35.811
24.9603 | -35.6271 | -35.8168
24.9664 | -35.6276 | -35.8225
24.9725 | -35.6279 | -35.8284
24.9786 | -35.6283 | -35.8337
24.9847 | -35.6287 | -35.8396
24.9908 | -35.6291 | -35.8454
24.9969 | -35.6295 | -35.8511
25.0031 | -35.6299 | -35.8571
25.0092 | -35.6303 | -35.863
25.0153 | -35.6306 | -35.8685
25.0214 | -35.631 | -35.8744
25.0275 | -35.6314 | -35.8802
25.0336 | -35.6318 | -35.886
25.0397 | -35.6322 | -35.8918
25.0458 | -35.6325 | -35.8974
25.0519 | -35.6329 | -35.9033
25.058 | -35.6333 | -35.9091
25.0641 | -35.6337 | -35.9148
25.0702 | -35.6341 | -35.9204
25.0763 | -35.6345 | -35.9261
25.0824 | -35.6348 | -35.9313
25.0885 | -35.6351 | -35.9371
25.0946 | -35.6355 | -35.9429
25.1007 | -35.6359 | -35.9489
25.1068 | -35.6363 | -35.9546
25.1129 | -35.6367 | -35.96
25.119 | -35.6371 | -35.9654
25.1251 | -35.6374 | -35.9712
25.1312 | -35.6378 | -35.977
25.1373 | -35.6382 | -35.9828
25.1435 | -35.6386 | -35.9886
25.1496 | -35.6389 | -35.9941
25.1557 | -35.6393 | -35.9999
25.1618 | -35.6397 | -36.0056
25.1679 | -35.6401 | -36.0114
25.174 | -35.6404 | -36.0172
25.1801 | -35.6408 | -36.023
25.1862 | -35.6411 | -36.0284
25.1923 | -35.6415 | -36.0342
25.1984 | -35.6419 | -36.0399
25.2045 | -35.6422 | -36.0455
25.2106 | -35.6426 | -36.0513
25.2167 | -35.643 | -36.0572
25.2228 | -35.6433 | -36.0628
25.2289 | -35.6437 | -36.0685
25.235 | -35.6441 | -36.0744
25.2411 | -35.6445 | -36.0801
25.2472 | -35.6449 | -36.0859
25.2533 | -35.6452 | -36.0916
25.2594 | -35.6455 | -36.0969
25.2655 | -35.6459 | -36.1025
25.2716 | -35.6463 | -36.108
25.2777 | -35.6466 | -36.1137
25.2838 | -35.647 | -36.1198
25.29 | -35.6474 | -36.1255
25.2961 | -35.6477 | -36.131
25.3022 | -35.6481 | -36.1365
25.3083 | -35.6485 | -36.1422
25.3144 | -35.6489 | -36.1481
25.3205 | -35.6492 | -36.1538
25.3266 | -35.6495 | -36.1594
25.3327 | -35.6499 | -36.1651
25.3388 | -35.6503 | -36.1707
25.3449 | -35.6507 | -36.1764
25.351 | -35.6511 | -36.1819
25.3571 | -35.6515 | -36.1877
25.3632 | -35.6518 | -36.1932
25.3693 | -35.6522 | -36.199
25.3754 | -35.6525 | -36.2046
25.3815 | -35.6529 | -36.2104
25.3876 | -35.6532 | -36.216
25.3937 | -35.6536 | -36.2217
25.3998 | -35.6539 | -36.227
25.4059 | -35.6543 | -36.2327
25.412 | -35.6546 | -36.2382
25.4181 | -35.655 | -36.2438
25.4242 | -35.6554 | -36.2494
25.4304 | -35.6558 | -36.255
25.4365 | -35.6561 | -36.2603
25.4426 | -35.6565 | -36.266
25.4487 | -35.6568 | -36.2718
25.4548 | -35.6572 | -36.2773
25.4609 | -35.6575 | -36.2829
25.467 | -35.6579 | -36.2883
25.4731 | -35.6582 | -36.2939
25.4792 | -35.6586 | -36.2996
25.4853 | -35.659 | -36.3054
25.4914 | -35.6593 | -36.3111
25.4975 | -35.6597 | -36.3166
25.5036 | -35.6601 | -36.3222
25.5097 | -35.6605 | -36.3279
25.5158 | -35.6608 | -36.3335
25.5219 | -35.6612 | -36.3394
25.528 | -35.6615 | -36.3451
25.5341 | -35.6619 | -36.3507
25.5402 | -35.6623 | -36.356
25.5463 | -35.6626 | -36.3615
25.5524 | -35.663 | -36.367
25.5585 | -35.6634 | -36.3725
25.5646 | -35.6637 | -36.3781
25.5707 | -35.6641 | -36.3838
25.5769 | -35.6644 | -36.3892
25.583 | -35.6648 | -36.3949
25.5891 | -35.6651 | -36.4005
25.5952 | -35.6654 | -36.4059
25.6013 | -35.6658 | -36.4116
25.6074 | -35.6661 | -36.4171
25.6135 | -35.6665 | -36.4225
25.6196 | -35.6668 | -36.4281
25.6257 | -35.6672 | -36.4334
25.6318 | -35.6675 | -36.4389
25.6379 | -35.6679 | -36.445
25.644 | -35.6683 | -36.4508
25.6501 | -35.6687 | -36.4563
25.6562 | -35.6691 | -36.4618
25.6623 | -35.6694 | -36.4672
25.6684 | -35.6698 | -36.4726
25.6745 | -35.6701 | -36.4782
25.6806 | -35.6705 | -36.4836
25.6867 | -35.6708 | -36.4893
25.6928 | -35.6712 | -36.4951
25.6989 | -35.6715 | -36.5006
25.705 | -35.6719 | -36.5064
25.7111 | -35.6722 | -36.5119
25.7173 | -35.6726 | -36.5172
25.7234 | -35.6729 | -36.523
25.7295 | -35.6733 | -36.5283
25.7356 | -35.6736 | -36.534
25.7417 | -35.6739 | -36.5395
25.7478 | -35.6742 | -36.5452
25.7539 | -35.6746 | -36.5507
25.76 | -35.6749 | -36.5561
25.7661 | -35.6752 | -36.5617
25.7722 | -35.6756 | -36.5675
25.7783 | -35.6759 | -36.573
25.7844 | -35.6762 | -36.5785
25.7905 | -35.6765 | -36.584
25.7966 | -35.6769 | -36.5895
25.8027 | -35.6772 | -36.595
25.8088 | -35.6776 | -36.6008
25.8149 | -35.678 | -36.6063
25.821 | -35.6783 | -36.6117
25.8271 | -35.6787 | -36.6173
25.8332 | -35.679 | -36.6228
25.8393 | -35.6794 | -36.6286
25.8454 | -35.6797 | -36.6343
25.8515 | -35.68 | -36.6398
25.8576 | -35.6804 | -36.6451
25.8638 | -35.6807 | -36.6508
25.8699 | -35.6811 | -36.6561
25.876 | -35.6814 | -36.6614
25.8821 | -35.6818 | -36.6672
25.8882 | -35.6821 | -36.6725
25.8943 | -35.6824 | -36.6781
25.9004 | -35.6827 | -36.6834
25.9065 | -35.6831 | -36.6889
25.9126 | -35.6835 | -36.6943
25.9187 | -35.6838 | -36.7001
25.9248 | -35.6842 | -36.7054
25.9309 | -35.6845 | -36.711
25.937 | -35.6848 | -36.7164
25.9431 | -35.6852 | -36.7219
25.9492 | -35.6854 | -36.7274
25.9553 | -35.6858 | -36.7329
25.9614 | -35.6861 | -36.7384
25.9675 | -35.6864 | -36.744
25.9736 | -35.6868 | -36.7495
25.9797 | -35.6871 | -36.7552
25.9858 | -35.6874 | -36.7605
25.9919 | -35.6877 | -36.7662
25.998 | -35.6881 | -36.7716
26.0042 | -35.6884 | -36.7772
26.0103 | -35.6888 | -36.7828
26.0164 | -35.6891 | -36.7882
26.0225 | -35.6894 | -36.7937
26.0286 | -35.6897 | -36.7992
26.0347 | -35.6901 | -36.8041
26.0408 | -35.6904 | -36.8098
26.0469 | -35.6908 | -36.8152
26.053 | -35.6911 | -36.8207
26.0591 | -35.6915 | -36.8262
26.0652 | -35.6918 | -36.8314
26.0713 | -35.6921 | -36.8369
26.0774 | -35.6924 | -36.8424
26.0835 | -35.6928 | -36.848
26.0896 | -35.6931 | -36.8534
26.0957 | -35.6934 | -36.8591
26.1018 | -35.6938 | -36.8643
26.1079 | -35.6941 | -36.8698
26.114 | -35.6944 | -36.8755
26.1201 | -35.6948 | -36.881
26.1262 | -35.6951 | -36.8865
26.1323 | -35.6954 | -36.8921
26.1384 | -35.6958 | -36.8975
26.1445 | -35.6962 | -36.903
26.1507 | -35.6965 | -36.9087
26.1568 | -35.6968 | -36.9144
26.1629 | -35.6972 | -36.9201
26.169 | -35.6975 | -36.9257
26.1751 | -35.6979 | -36.9311
26.1812 | -35.6983 | -36.9365
26.1873 | -35.6986 | -36.9419
26.1934 | -35.6989 | -36.9472
26.1995 | -35.6993 | -36.9525
26.2056 | -35.6996 | -36.9578
26.2117 | -35.7 | -36.9633
26.2178 | -35.7003 | -36.9687
26.2239 | -35.7006 | -36.9742
26.23 | -35.701 | -36.9797
26.2361 | -35.7013 | -36.9852
26.2422 | -35.7015 | -36.9904
26.2483 | -35.7019 | -36.9959
26.2544 | -35.7022 | -37.0015
26.2605 | -35.7025 | -37.007
26.2666 | -35.7029 | -37.0127
26.2727 | -35.7032 | -37.0184
26.2788 | -35.7036 | -37.0237
26.2849 | -35.7039 | -37.0293
26.2911 | -35.7042 | -37.0347
26.2972 | -35.7046 | -37.0403
26.3033 | -35.705 | -37.0457
26.3094 | -35.7053 | -37.0511
26.3155 | -35.7056 | -37.0567
26.3216 | -35.706 | -37.0623
26.3277 | -35.7063 | -37.0676
26.3338 | -35.7066 | -37.0733
26.3399 | -35.7069 | -37.0785
26.346 | -35.7072 | -37.0838
26.3521 | -35.7076 | -37.0896
26.3582 | -35.708 | -37.0951
26.3643 | -35.7083 | -37.1008
26.3704 | -35.7086 | -37.1064
26.3765 | -35.7089 | -37.112
26.3826 | -35.7092 | -37.117
26.3887 | -35.7096 | -37.1227
26.3948 | -35.7099 | -37.1283
26.4009 | -35.7103 | -37.1338
26.407 | -35.7106 | -37.139
26.4131 | -35.711 | -37.1442
26.4192 | -35.7113 | -37.1495
26.4253 | -35.7116 | -37.1551
26.4314 | -35.712 | -37.1606
26.4376 | -35.7123 | -37.166
26.4437 | -35.7127 | -37.1716
26.4498 | -35.713 | -37.1773
26.4559 | -35.7133 | -37.1827
26.462 | -35.7136 | -37.1884
26.4681 | -35.714 | -37.1939
26.4742 | -35.7143 | -37.1993
26.4803 | -35.7147 | -37.2051
26.4864 | -35.715 | -37.2102
26.4925 | -35.7154 | -37.2156
26.4986 | -35.7157 | -37.2211
26.5047 | -35.7161 | -37.2267
26.5108 | -35.7164 | -37.2324
26.5169 | -35.7167 | -37.2379
26.523 | -35.717 | -37.243
26.5291 | -35.7174 | -37.2484
26.5352 | -35.7177 | -37.2541
26.5413 | -35.718 | -37.2596
26.5474 | -35.7183 | -37.265
26.5535 | -35.7187 | -37.2704
26.5596 | -35.719 | -37.2757
26.5657 | -35.7193 | -37.2811
26.5718 | -35.7196 | -37.2868
26.578 | -35.7199 | -37.2921
26.5841 | -35.7203 | -37.2974
26.5902 | -35.7206 | -37.3028
26.5963 | -35.721 | -37.308
26.6024 | -35.7213 | -37.3134
26.6085 | -35.7217 | -37.3187
26.6146 | -35.7219 | -37.3241
26.6207 | -35.7222 | -37.3296
26.6268 | -35.7226 | -37.335
26.6329 | -35.7229 | -37.3404
26.639 | -35.7232 | -37.3457
26.6451 | -35.7235 | -37.3513
26.6512 | -35.7238 | -37.3565
26.6573 | -35.7241 | -37.3617
26.6634 | -35.7245 | -37.3675
26.6695 | -35.7248 | -37.3728
26.6756 | -35.7251 | -37.3782
26.6817 | -35.7254 | -37.3839
26.6878 | -35.7258 | -37.3894
26.6939 | -35.7261 | -37.3948
26.7 | -35.7264 | -37.3997
26.7061 | -35.7267 | -37.4051
26.7122 | -35.727 | -37.4107
26.7183 | -35.7274 | -37.416
26.7245 | -35.7277 | -37.4216
26.7306 | -35.728 | -37.4271
26.7367 | -35.7283 | -37.432
26.7428 | -35.7287 | -37.4375
26.7489 | -35.729 | -37.4428
26.755 | -35.7292 | -37.448
26.7611 | -35.7296 | -37.4536
26.7672 | -35.7299 | -37.4588
26.7733 | -35.7302 | -37.4639
26.7794 | -35.7306 | -37.4695
26.7855 | -35.7309 | -37.4749
26.7916 | -35.7312 | -37.4804
26.7977 | -35.7315 | -37.4858
26.8038 | -35.7318 | -37.4911
26.8099 | -35.7321 | -37.4964
26.816 | -35.7324 | -37.5016
26.8221 | -35.7328 | -37.5069
26.8282 | -35.733 | -37.512
26.8343 | -35.7334 | -37.5174
26.8404 | -35.7337 | -37.5228
26.8465 | -35.734 | -37.528
26.8526 | -35.7343 | -37.5334
26.8587 | -35.7346 | -37.5387
26.8649 | -35.7349 | -37.544
26.871 | -35.7352 | -37.5493
26.8771 | -35.7356 | -37.5542
26.8832 | -35.736 | -37.5597
26.8893 | -35.7363 | -37.5651
26.8954 | -35.7366 | -37.5705
26.9015 | -35.7369 | -37.5762
26.9076 | -35.7372 | -37.5816
26.9137 | -35.7375 | -37.5868
26.9198 | -35.7378 | -37.5921
26.9259 | -35.7381 | -37.5979
26.932 | -35.7385 | -37.603
26.9381 | -35.7387 | -37.6083
26.9442 | -35.739 | -37.6134
26.9503 | -35.7393 | -37.6187
26.9564 | -35.7397 | -37.624
26.9625 | -35.74 | -37.6293
26.9686 | -35.7404 | -37.6349
26.9747 | -35.7407 | -37.6401
26.9808 | -35.7411 | -37.6452
26.9869 | -35.7414 | -37.6503
26.993 | -35.7417 | -37.656
26.9991 | -35.742 | -37.6611
27.0052 | -35.7424 | -37.6661
27.0114 | -35.7427 | -37.6717
27.0175 | -35.743 | -37.6764
27.0236 | -35.7433 | -37.6818
27.0297 | -35.7436 | -37.687
27.0358 | -35.7439 | -37.6923
27.0419 | -35.7443 | -37.6978
27.048 | -35.7446 | -37.7032
27.0541 | -35.7449 | -37.7084
27.0602 | -35.7452 | -37.7134
27.0663 | -35.7455 | -37.719
27.0724 | -35.7459 | -37.7243
27.0785 | -35.7462 | -37.7297
27.0846 | -35.7465 | -37.7349
27.0907 | -35.7469 | -37.7404
27.0968 | -35.7472 | -37.7459
27.1029 | -35.7475 | -37.7511
27.109 | -35.7478 | -37.7565
27.1151 | -35.7482 | -37.762
27.1212 | -35.7484 | -37.7671
27.1273 | -35.7488 | -37.772
27.1334 | -35.7491 | -37.7774
27.1395 | -35.7494 | -37.7827
27.1456 | -35.7497 | -37.788
27.1518 | -35.7501 | -37.7934
27.1579 | -35.7504 | -37.7989
27.164 | -35.7507 | -37.8041
27.1701 | -35.751 | -37.8097
27.1762 | -35.7512 | -37.815
27.1823 | -35.7515 | -37.8199
27.1884 | -35.7518 | -37.8253
27.1945 | -35.7521 | -37.8306
27.2006 | -35.7525 | -37.8359
27.2067 | -35.7527 | -37.8412
27.2128 | -35.7531 | -37.8464
27.2189 | -35.7534 | -37.8518
27.225 | -35.7537 | -37.8575
27.2311 | -35.754 | -37.8631
27.2372 | -35.7544 | -37.8686
27.2433 | -35.7547 | -37.8738
27.2494 | -35.755 | -37.8792
27.2555 | -35.7553 | -37.8847
27.2616 | -35.7556 | -37.8896
27.2677 | -35.7559 | -37.8951
27.2738 | -35.7562 | -37.9001
27.2799 | -35.7565 | -37.9052
27.286 | -35.7568 | -37.9102
27.2921 | -35.7571 | -37.9156
27.2983 | -35.7574 | -37.9208
27.3044 | -35.7577 | -37.9262
27.3105 | -35.758 | -37.9314
27.3166 | -35.7583 | -37.9366
27.3227 | -35.7587 | -37.9419
27.3288 | -35.7589 | -37.9473
27.3349 | -35.7593 | -37.9523
27.341 | -35.7596 | -37.9573
27.3471 | -35.7599 | -37.9628
27.3532 | -35.7602 | -37.9681
27.3593 | -35.7605 | -37.9732
27.3654 | -35.7607 | -37.9784
27.3715 | -35.761 | -37.9837
27.3776 | -35.7613 | -37.989
27.3837 | -35.7617 | -37.994
27.3898 | -35.762 | -37.9994
27.3959 | -35.7623 | -38.0046
27.402 | -35.7626 | -38.0098
27.4081 | -35.7629 | -38.015
27.4142 | -35.7632 | -38.0203
27.4203 | -35.7635 | -38.0258
27.4264 | -35.7638 | -38.0312
27.4325 | -35.7641 | -38.0365
27.4387 | -35.7644 | -38.0419
27.4448 | -35.7647 | -38.0473
27.4509 | -35.7651 | -38.0525
27.457 | -35.7654 | -38.0579
27.4631 | -35.7657 | -38.0629
27.4692 | -35.7659 | -38.0682
27.4753 | -35.7662 | -38.0734
27.4814 | -35.7665 | -38.0787
27.4875 | -35.7668 | -38.0837
27.4936 | -35.7672 | -38.089
27.4997 | -35.7675 | -38.0944
27.5058 | -35.7678 | -38.0998
27.5119 | -35.768 | -38.1053
27.518 | -35.7683 | -38.1107
27.5241 | -35.7686 | -38.1157
27.5302 | -35.7689 | -38.121
27.5363 | -35.7693 | -38.126
27.5424 | -35.7696 | -38.1309
27.5485 | -35.77 | -38.136
27.5546 | -35.7703 | -38.1415
27.5607 | -35.7706 | -38.1464
27.5668 | -35.7709 | -38.1516
27.5729 | -35.7712 | -38.1566
27.5791 | -35.7714 | -38.1613
27.5852 | -35.7717 | -38.1666
27.5913 | -35.7721 | -38.172
27.5974 | -35.7724 | -38.1774
27.6035 | -35.7727 | -38.183
27.6096 | -35.773 | -38.1883
27.6157 | -35.7733 | -38.1937
27.6218 | -35.7736 | -38.1989
27.6279 | -35.7739 | -38.2042
27.634 | -35.7742 | -38.2093
27.6401 | -35.7745 | -38.2146
27.6462 | -35.7748 | -38.2196
27.6523 | -35.7751 | -38.2245
27.6584 | -35.7754 | -38.2298
27.6645 | -35.7757 | -38.2351
27.6706 | -35.776 | -38.2403
27.6767 | -35.7764 | -38.2456
27.6828 | -35.7766 | -38.2506
27.6889 | -35.777 | -38.2559
27.695 | -35.7773 | -38.2612
27.7011 | -35.7776 | -38.2663
27.7072 | -35.7779 | -38.2715
27.7133 | -35.7782 | -38.2767
27.7194 | -35.7785 | -38.2816
27.7256 | -35.7788 | -38.287
27.7317 | -35.7791 | -38.2921
27.7378 | -35.7794 | -38.2972
27.7439 | -35.7796 | -38.3025
27.75 | -35.7799 | -38.3077
27.7561 | -35.7802 | -38.3124
27.7622 | -35.7805 | -38.3176
27.7683 | -35.7808 | -38.3226
27.7744 | -35.7811 | -38.3281
27.7805 | -35.7814 | -38.3331
27.7866 | -35.7816 | -38.3385
27.7927 | -35.7819 | -38.3433
27.7988 | -35.7822 | -38.3486
27.8049 | -35.7825 | -38.3541
27.811 | -35.7828 | -38.359
27.8171 | -35.783 | -38.3643
27.8232 | -35.7834 | -38.3689
27.8293 | -35.7836 | -38.3743
27.8354 | -35.7839 | -38.3793
27.8415 | -35.7843 | -38.3845
27.8476 | -35.7846 | -38.3898
27.8537 | -35.7849 | -38.395
27.8598 | -35.7852 | -38.3999
27.866 | -35.7855 | -38.4053
27.8721 | -35.7858 | -38.4105
27.8782 | -35.786 | -38.4158
27.8843 | -35.7863 | -38.4207
27.8904 | -35.7866 | -38.4258
27.8965 | -35.7869 | -38.4308
27.9026 | -35.7872 | -38.4359
27.9087 | -35.7876 | -38.4411
27.9148 | -35.7879 | -38.4464
27.9209 | -35.7882 | -38.4517
27.927 | -35.7884 | -38.4566
27.9331 | -35.7887 | -38.4618
27.9392 | -35.789 | -38.467
27.9453 | -35.7893 | -38.4725
27.9514 | -35.7896 | -38.4777
27.9575 | -35.7899 | -38.4828
27.9636 | -35.7902 | -38.4882
27.9697 | -35.7905 | -38.4931
27.9758 | -35.7908 | -38.4981
27.9819 | -35.791 | -38.5036
27.988 | -35.7913 | -38.5087
27.9941 | -35.7916 | -38.5139
28.0002 | -35.7919 | -38.5189
28.0063 | -35.7922 | -38.5242
28.0125 | -35.7925 | -38.5295
28.0186 | -35.7927 | -38.5349
28.0247 | -35.793 | -38.5395
28.0308 | -35.7933 | -38.545
28.0369 | -35.7936 | -38.5501
28.043 | -35.7938 | -38.5554
28.0491 | -35.7941 | -38.5606
28.0552 | -35.7944 | -38.5656
28.0613 | -35.7947 | -38.5709
28.0674 | -35.795 | -38.5761
28.0735 | -35.7952 | -38.5812

## ACOUS_TF
Frequency(Hz) | Amplitude (dBV/Pa) | Phase (degrees)
28 | -36.5143 | -43.4101
21.4908 | -35.5276 | -35.5796
16.4948 | -34.9112 | -28.0972
12.6602 | -34.549 | -21.7483
9.7171 | -34.3295 | -16.7812
7.4582 | -34.1832 | -12.7531
5.7244 | -34.0739 | -9.4894
4.3936 | -34.004 | -6.8583
3.3722 | -33.9892 | -4.7812
2.5883 | -33.9723 | -3.2065
1.9866 | -33.9704 | -1.9486
1.5248 | -33.9755 | -0.94283
1.1703 | -33.9791 | -0.28253
1 | -33.9817 | 0.043294
0.89822 | -33.9857 | 0.3177
0.68941 | -33.9656 | 0.90857
0.52919 | -33.9607 | 1.3666
0.40612 | -33.9657 | 1.977
0.31171 | -33.9468 | 2.4356
0.23928 | -33.9674 | 3.1591
0.18358 | -33.9478 | 3.5863
0.14091 | -33.9569 | 4.4782
0.10819 | -33.9834 | 5.64
0.083017 | -34.0111 | 7.1247
0.063705 | -34.0485 | 9.006
0.048923 | -34.1073 | 11.5149
0.037479 | -34.2079 | 14.7118
0.028801 | -34.3991 | 18.7377
0.022078 | -34.7081 | 24.022
0.017023 | -35.1787 | 30.4114
0.013018 | -35.9802 | 38.1225
0.010014 | -37.0972 | 46.0424

## SELF-NOISE
Frequency (Hz) | PSD (dBPa/\/Hz)
0.003759 | -16.8245
0.007517 | -35.3455
0.011276 | -40.2214
0.015034 | -42.9189
0.018793 | -45.4943
0.022551 | -47.378
0.02631 | -48.7028
0.030068 | -49.8867
0.033827 | -50.99
0.037585 | -51.9305
0.041344 | -53.3089
0.045102 | -55.037
0.048861 | -56.4558
0.052619 | -56.1387
0.056378 | -56.1287
0.060136 | -57.0643
0.063895 | -58.7295
0.067653 | -60.0784
0.071412 | -58.8247
0.07517 | -60.3322
0.078929 | -61.2631
0.082687 | -60.89
0.086446 | -60.4794
0.090204 | -62.4697
0.093963 | -64.2657
0.097721 | -63.4197
0.10148 | -63.0041
0.10524 | -63.5597
0.109 | -64.443
0.11275 | -64.1428
0.11651 | -64.1751
0.12027 | -64.9414
0.12403 | -66.2997
0.12779 | -67.6644
0.13155 | -67.5355
0.13531 | -67.2498
0.13906 | -66.5546
0.14282 | -67.3015
0.14658 | -67.8312
0.15034 | -68.1734
0.1541 | -68.8731
0.15786 | -68.4813
0.16162 | -68.3291
0.16537 | -67.6697
0.16913 | -69.33
0.17289 | -70.5611
0.17665 | -70.2613
0.18041 | -69.5371
0.18417 | -70.2376
0.18793 | -70.5009
0.19168 | -70.7045
0.19544 | -70.9471
0.1992 | -71.679
0.20296 | -71.8887
0.20672 | -71.7524
0.21048 | -71.438
0.21424 | -72.0772
0.21799 | -72.413
0.22175 | -71.2812
0.22551 | -72.5123
0.22927 | -73.2359
0.23303 | -73.3006
0.23679 | -72.7144
0.24055 | -73.0144
0.2443 | -73.0929
0.24806 | -73.3522
0.25182 | -73.4124
0.25558 | -73.8693
0.25934 | -73.8083
0.2631 | -73.1324
0.26686 | -73.032
0.27061 | -73.8516
0.27437 | -74.2115
0.27813 | -74.9312
0.28189 | -75.2261
0.28565 | -74.6531
0.28941 | -74.4413
0.29316 | -73.9483
0.29692 | -75.016
0.30068 | -75.0097
0.30444 | -74.7938
0.3082 | -74.3256
0.31196 | -75.8206
0.31572 | -76.4989
0.31947 | -76.0778
0.32323 | -76.2189
0.32699 | -76.938
0.33075 | -76.7853
0.33451 | -76.6379
0.33827 | -76.6644
0.34203 | -76.3658
0.34578 | -76.5253
0.34954 | -76.3034
0.3533 | -75.5819
0.35706 | -75.8959
0.36082 | -76.7795
0.36458 | -76.9311
0.36834 | -77.804
0.37209 | -77.9519
0.37585 | -77.7696
0.37961 | -77.6005
0.38337 | -77.3342
0.38713 | -77.5111
0.39089 | -77.8138
0.39464 | -77.0303
0.3984 | -76.6256
0.40216 | -76.7518
0.40592 | -77.1749
0.40968 | -77.7042
0.41344 | -78.1561
0.41719 | -78.2599
0.42095 | -78.9764
0.42471 | -79.3371
0.42847 | -78.9469
0.43223 | -77.731
0.43599 | -78.1316
0.43975 | -78.4472
0.4435 | -78.5298
0.44726 | -78.4446
0.45102 | -79.2114
0.45478 | -79.1704
0.45854 | -79.0439
0.4623 | -79.3981
0.46606 | -79.8132
0.46981 | -79.3241
0.47357 | -79.4772
0.47733 | -80.2003
0.48109 | -79.0003
0.48485 | -78.8077
0.48861 | -79.4879
0.49237 | -79.5984
0.49612 | -79.6177
0.49988 | -79.6889
0.50364 | -79.1468
0.5074 | -79.9042
0.51116 | -80.7887
0.51492 | -81.5147
0.51867 | -81.8539
0.52243 | -80.1935
0.52619 | -81.7084
0.52995 | -81.1408
0.53371 | -80.5719
0.53747 | -80.62
0.54123 | -80.4376
0.54499 | -80.7709
0.54874 | -80.9475
0.5525 | -80.8415
0.55626 | -80.8989
0.56002 | -80.7157
0.56378 | -80.7563
0.56754 | -81.2095
0.57129 | -81.2127
0.57505 | -80.9732
0.57881 | -81.0909
0.58257 | -81.6144
0.58633 | -82.022
0.59009 | -82.2497
0.59384 | -82.1567
0.5976 | -82.2245
0.60136 | -82.7701
0.60512 | -83.0891
0.60888 | -83.0984
0.61264 | -82.7904
0.6164 | -82.3681
0.62016 | -82.1862
0.62391 | -81.8817
0.62767 | -81.5568
0.63143 | -81.5297
0.63519 | -81.6302
0.63895 | -81.9376
0.64271 | -81.9391
0.64646 | -81.7351
0.65022 | -81.9115
0.65398 | -82.4202
0.65774 | -82.8863
0.6615 | -82.8353
0.66526 | -82.8419
0.66902 | -82.7815
0.67277 | -82.4549
0.67653 | -82.4649
0.68029 | -82.5911
0.68405 | -82.7602
0.68781 | -83.4111
0.69157 | -84.2998
0.69532 | -84.5196
0.69908 | -83.9265
0.70284 | -83.4284
0.7066 | -83.643
0.71036 | -84.1872
0.71412 | -84.2222
0.71788 | -84.1619
0.72164 | -84.1704
0.72539 | -84.1224
0.72915 | -84.1695
0.73291 | -84.5109
0.73667 | -84.8277
0.74043 | -84.5122
0.74419 | -84.0957
0.74794 | -84.0703
0.7517 | -83.9865
0.75546 | -83.7683
0.75922 | -83.2699
0.76298 | -82.9943
0.76674 | -83.4712
0.7705 | -84.1445
0.77425 | -84.2861
0.77801 | -84.4427
0.78177 | -84.8791
0.78553 | -84.9629
0.78929 | -84.5711
0.79305 | -84.4319
0.7968 | -84.5609
0.80056 | -84.81
0.80432 | -84.9791
0.80808 | -84.9996
0.81184 | -85.3587
0.8156 | -85.5545
0.81936 | -85.5185
0.82312 | -85.6645
0.82687 | -85.4592
0.83063 | -85.0489
0.83439 | -85.0339
0.83815 | -85.1693
0.84191 | -84.9476
0.84567 | -84.893
0.84942 | -85.4399
0.85318 | -86.3592
0.85694 | -86.543
0.8607 | -85.6599
0.86446 | -85.1367
0.86822 | -85.5396
0.87198 | -86.0003
0.87573 | -85.9945
0.87949 | -86.0814
0.88325 | -86.412
0.88701 | -86.5029
0.89077 | -86.3149
0.89453 | -86.3267
0.89829 | -86.5301
0.90204 | -86.5856
0.9058 | -86.2683
0.90956 | -85.6979
0.91332 | -85.0138
0.91708 | -84.6915
0.92084 | -85.0031
0.92459 | -85.5225
0.92835 | -85.6438
0.93211 | -85.5206
0.93587 | -85.5002
0.93963 | -85.6724
0.94339 | -85.7829
0.94715 | -85.7708
0.9509 | -86.0962
0.95466 | -86.5714
0.95842 | -86.7632
0.96218 | -86.783
0.96594 | -86.7577
0.9697 | -86.5782
0.97345 | -86.2286
0.97721 | -86.1834
0.98097 | -86.5687
0.98473 | -86.6311
0.98849 | -86.2911
0.99225 | -86.3052
0.99601 | -86.3862
0.99977 | -86.6855
1.0035 | -87.2787
1.0073 | -87.5842
1.011 | -87.39
1.0148 | -87.0536
1.0186 | -86.7869
1.0223 | -86.7202
1.0261 | -86.5216
1.0298 | -86.4894
1.0336 | -86.6247
1.0373 | -86.9837
1.0411 | -87.2593
1.0449 | -87.4194
1.0486 | -87.6099
1.0524 | -87.5133
1.0561 | -87.2972
1.0599 | -87.3718
1.0637 | -87.3945
1.0674 | -87.1676
1.0712 | -87.3285
1.0749 | -87.8087
1.0787 | -88.0965
1.0825 | -88.0444
1.0862 | -87.9378
1.09 | -87.685
1.0937 | -87.3095
1.0975 | -86.9884
1.1012 | -86.885
1.105 | -87.0265
1.1088 | -87.3977
1.1125 | -87.8166
1.1163 | -88.0012
1.12 | -87.9787
1.1238 | -87.898
1.1276 | -87.8102
1.1313 | -87.7914
1.1351 | -87.8801
1.1388 | -88.017
1.1426 | -88.0968
1.1463 | -88.1815
1.1501 | -88.3898
1.1539 | -88.5309
1.1576 | -88.4781
1.1614 | -88.4249
1.1651 | -88.5168
1.1689 | -88.7252
1.1727 | -88.9133
1.1764 | -88.9668
1.1802 | -88.8959
1.1839 | -88.8077
1.1877 | -88.8287
1.1914 | -88.965
1.1952 | -89.0589
1.199 | -88.9755
1.2027 | -88.7692
1.2065 | -88.6103
1.2102 | -88.6708
1.214 | -88.9971
1.2178 | -89.4471
1.2215 | -89.7576
1.2253 | -89.7738
1.229 | -89.7046
1.2328 | -89.6812
1.2366 | -89.4686
1.2403 | -89.0338
1.2441 | -88.6823
1.2478 | -88.6044
1.2516 | -88.7008
1.2553 | -88.7132
1.2591 | -88.6178
1.2629 | -88.5618
1.2666 | -88.4933
1.2704 | -88.3649
1.2741 | -88.2771
1.2779 | -88.2894
1.2817 | -88.4117
1.2854 | -88.691
1.2892 | -89.0683
1.2929 | -89.4412
1.2967 | -89.808
1.3004 | -90.1339
1.3042 | -90.2822
1.308 | -90.1694
1.3117 | -89.9337
1.3155 | -89.7894
1.3192 | -89.7591
1.323 | -89.8351
1.3268 | -89.9813
1.3305 | -89.9843
1.3343 | -89.7478
1.338 | -89.509
1.3418 | -89.4386
1.3455 | -89.5431
1.3493 | -89.7287
1.3531 | -89.8312
1.3568 | -89.8712
1.3606 | -89.9573
1.3643 | -90.0638
1.3681 | -90.2368
1.3719 | -90.4955
1.3756 | -90.6762
1.3794 | -90.6465
1.3831 | -90.5041
1.3869 | -90.3819
1.3907 | -90.3824
1.3944 | -90.5645
1.3982 | -90.8476
1.4019 | -90.9422
1.4057 | -90.6452
1.4094 | -90.2283
1.4132 | -90.02
1.417 | -90.0179
1.4207 | -90.0884
1.4245 | -90.2363
1.4282 | -90.6317
1.432 | -91.1798
1.4358 | -91.1743
1.4395 | -90.4043
1.4433 | -89.7369
1.447 | -89.7125
1.4508 | -90.1286
1.4545 | -90.4415
1.4583 | -90.4741
1.4621 | -90.4994
1.4658 | -90.6788
1.4696 | -90.8871
1.4733 | -90.9147
1.4771 | -90.7812
1.4809 | -90.7204
1.4846 | -90.7793
1.4884 | -90.6836
1.4921 | -90.3452
1.4959 | -90.0846
1.4996 | -90.1108
1.5034 | -90.2947
1.5072 | -90.4151
1.5109 | -90.3633
1.5147 | -90.1409
1.5184 | -89.8676
1.5222 | -89.7983
1.526 | -90.0658
1.5297 | -90.4673
1.5335 | -90.6322
1.5372 | -90.5362
1.541 | -90.491
1.5448 | -90.7235
1.5485 | -91.102
1.5523 | -91.2601
1.556 | -91.0898
1.5598 | -90.8938
1.5635 | -90.8181
1.5673 | -90.7528
1.5711 | -90.6339
1.5748 | -90.5321
1.5786 | -90.4956
1.5823 | -90.5201
1.5861 | -90.677
1.5899 | -90.924
1.5936 | -90.9985
1.5974 | -90.8693
1.6011 | -90.7351
1.6049 | -90.7211
1.6086 | -90.7566
1.6124 | -90.7945
1.6162 | -90.9754
1.6199 | -91.2486
1.6237 | -91.3625
1.6274 | -91.4762
1.6312 | -91.5296
1.635 | -91.4292
1.6387 | -91.1812
1.6425 | -90.94
1.6462 | -90.8429
1.65 | -90.932
1.6537 | -91.176
1.6575 | -91.475
1.6613 | -91.7114
1.665 | -91.8319
1.6688 | -91.8843
1.6725 | -91.9186
1.6763 | -91.9333
1.6801 | -91.9356
1.6838 | -91.9472
1.6876 | -92.0013
1.6913 | -92.0648
1.6951 | -92.0419
1.6988 | -91.8926
1.7026 | -91.6891
1.7064 | -91.5416
1.7101 | -91.4818
1.7139 | -91.4717
1.7176 | -91.4997
1.7214 | -91.5739
1.7252 | -91.6977
1.7289 | -91.8069
1.7327 | -91.8106
1.7364 | -91.7029
1.7402 | -91.5909
1.744 | -91.5907
1.7477 | -91.7108
1.7515 | -91.8892
1.7552 | -92.0535
1.759 | -92.1536
1.7627 | -92.1929
1.7665 | -92.1909
1.7703 | -92.1357
1.774 | -92.0271
1.7778 | -91.9091
1.7815 | -91.865
1.7853 | -91.893
1.7891 | -91.9401
1.7928 | -92.0182
1.7966 | -92.177
1.8003 | -92.3821
1.8041 | -92.4855
1.8078 | -92.4254
1.8116 | -92.3094
1.8154 | -92.257
1.8191 | -92.2912
1.8229 | -92.3421
1.8266 | -92.2888
1.8304 | -92.0969
1.8342 | -91.874
1.8379 | -91.7742
1.8417 | -91.8522
1.8454 | -92.0686
1.8492 | -92.3022
1.8529 | -92.3756
1.8567 | -92.2239
1.8605 | -91.9817
1.8642 | -91.7888
1.868 | -91.6848
1.8717 | -91.6402
1.8755 | -91.6574
1.8793 | -91.7591
1.883 | -91.9358
1.8868 | -92.0705
1.8905 | -92.0428
1.8943 | -91.8967
1.8981 | -91.7883
1.9018 | -91.8216
1.9056 | -92.0354
1.9093 | -92.3941
1.9131 | -92.7982
1.9168 | -93.0959
1.9206 | -93.169
1.9244 | -93.0209
1.9281 | -92.8038
1.9319 | -92.6741
1.9356 | -92.6777
1.9394 | -92.7473
1.9432 | -92.8268
1.9469 | -92.9097
1.9507 | -92.9817
1.9544 | -93.0001
1.9582 | -92.9792
1.9619 | -93.0147
1.9657 | -93.1534
1.9695 | -93.337
1.9732 | -93.4579
1.977 | -93.4314
1.9807 | -93.2787
1.9845 | -93.078
1.9883 | -92.9379
1.992 | -92.9299
1.9958 | -93.0438
1.9995 | -93.1827
2.0033 | -93.236
2.007 | -93.1825
2.0108 | -93.0727
2.0146 | -92.986
2.0183 | -93.0142
2.0221 | -93.1575
2.0258 | -93.3056
2.0296 | -93.333
2.0334 | -93.2568
2.0371 | -93.2028
2.0409 | -93.2426
2.0446 | -93.3718
2.0484 | -93.5381
2.0521 | -93.6499
2.0559 | -93.623
2.0597 | -93.4586
2.0634 | -93.2508
2.0672 | -93.116
2.0709 | -93.1373
2.0747 | -93.3378
2.0785 | -93.6278
2.0822 | -93.9098
2.086 | -94.1152
2.0897 | -94.201
2.0935 | -94.1932
2.0973 | -94.137
2.101 | -94.0566
2.1048 | -93.9472
2.1085 | -93.8116
2.1123 | -93.6519
2.116 | -93.4865
2.1198 | -93.3544
2.1236 | -93.2653
2.1273 | -93.1899
2.1311 | -93.1416
2.1348 | -93.1486
2.1386 | -93.2083
2.1424 | -93.3279
2.1461 | -93.5151
2.1499 | -93.7171
2.1536 | -93.8643
2.1574 | -93.9169
2.1611 | -93.8739
2.1649 | -93.7789
2.1687 | -93.7415
2.1724 | -93.727
2.1762 | -93.7487
2.1799 | -93.7903
2.1837 | -93.8053
2.1875 | -93.7771
2.1912 | -93.7157
2.195 | -93.6654
2.1987 | -93.6645
2.2025 | -93.7027
2.2062 | -93.7431
2.21 | -93.7562
2.2138 | -93.7383
2.2175 | -93.7105
2.2213 | -93.6843
2.225 | -93.6939
2.2288 | -93.7248
2.2326 | -93.7497
2.2363 | -93.7155
2.2401 | -93.5867
2.2438 | -93.3757
2.2476 | -93.1485
2.2514 | -92.9858
2.2551 | -92.9544
2.2589 | -93.0627
2.2626 | -93.263
2.2664 | -93.4951
2.2701 | -93.7003
2.2739 | -93.8358
2.2777 | -93.9095
2.2814 | -93.9506
2.2852 | -93.9779
2.2889 | -93.984
2.2927 | -93.928
2.2965 | -93.8072
2.3002 | -93.683
2.304 | -93.5975
2.3077 | -93.5631
2.3115 | -93.5696
2.3152 | -93.6056
2.319 | -93.6606
2.3228 | -93.7155
2.3265 | -93.79
2.3303 | -93.9025
2.334 | -94.072
2.3378 | -94.2627
2.3416 | -94.4148
2.3453 | -94.508
2.3491 | -94.53
2.3528 | -94.5031
2.3566 | -94.485
2.3603 | -94.5219
2.3641 | -94.6369
2.3679 | -94.7778
2.3716 | -94.9078
2.3754 | -94.9859
2.3791 | -95.009
2.3829 | -94.9565
2.3867 | -94.8105
2.3904 | -94.5918
2.3942 | -94.3487
2.3979 | -94.1184
2.4017 | -93.96
2.4055 | -93.9173
2.4092 | -93.9756
2.413 | -94.0864
2.4167 | -94.1939
2.4205 | -94.2561
2.4242 | -94.2764
2.428 | -94.289
2.4318 | -94.3357
2.4355 | -94.4144
2.4393 | -94.5005
2.443 | -94.5413
2.4468 | -94.514
2.4506 | -94.4592
2.4543 | -94.4128
2.4581 | -94.4037
2.4618 | -94.412
2.4656 | -94.43
2.4693 | -94.4423
2.4731 | -94.4655
2.4769 | -94.4961
2.4806 | -94.5099
2.4844 | -94.4812
2.4881 | -94.4036
2.4919 | -94.3168
2.4957 | -94.2563
2.4994 | -94.2453
2.5032 | -94.3145
2.5069 | -94.4602
2.5107 | -94.6508
2.5144 | -94.8171
2.5182 | -94.9292
2.522 | -94.9813
2.5257 | -95.0036
2.5295 | -95.0285
2.5332 | -95.0391
2.537 | -95.018
2.5408 | -94.9415
2.5445 | -94.8134
2.5483 | -94.6633
2.552 | -94.5215
2.5558 | -94.398
2.5595 | -94.3281
2.5633 | -94.3551
2.5671 | -94.4889
2.5708 | -94.6993
2.5746 | -94.9585
2.5783 | -95.2283
2.5821 | -95.4501
2.5859 | -95.5518
2.5896 | -95.5178
2.5934 | -95.4264
2.5971 | -95.3836
2.6009 | -95.4068
2.6047 | -95.4602
2.6084 | -95.4947
2.6122 | -95.4955
2.6159 | -95.4638
2.6197 | -95.3924
2.6234 | -95.2839
2.6272 | -95.1781
2.631 | -95.111
2.6347 | -95.0843
2.6385 | -95.0659
2.6422 | -95.0358
2.646 | -95.0024
2.6498 | -94.9901
2.6535 | -95.0185
2.6573 | -95.0715
2.661 | -95.1287
2.6648 | -95.1537
2.6685 | -95.1283
2.6723 | -95.0843
2.6761 | -95.0837
2.6798 | -95.1675
2.6836 | -95.3301
2.6873 | -95.5357
2.6911 | -95.7343
2.6949 | -95.8707
2.6986 | -95.9369
2.7024 | -95.9478
2.7061 | -95.9585
2.7099 | -95.9785
2.7136 | -95.9882
2.7174 | -95.9821
2.7212 | -95.9542
2.7249 | -95.9367
2.7287 | -95.9423
2.7324 | -95.9478
2.7362 | -95.9488
2.74 | -95.9417
2.7437 | -95.9304
2.7475 | -95.9261
2.7512 | -95.9175
2.755 | -95.901
2.7588 | -95.8818
2.7625 | -95.8568
2.7663 | -95.8264
2.77 | -95.7755
2.7738 | -95.6894
2.7775 | -95.5514
2.7813 | -95.3789
2.7851 | -95.2022
2.7888 | -95.0946
2.7926 | -95.0872
2.7963 | -95.1738
2.8001 | -95.3295
2.8039 | -95.5138
2.8076 | -95.6995
2.8114 | -95.8823
2.8151 | -96.049
2.8189 | -96.1582
2.8226 | -96.185
2.8264 | -96.1709
2.8302 | -96.1556
2.8339 | -96.1801
2.8377 | -96.2475
2.8414 | -96.3631
2.8452 | -96.5026
2.849 | -96.6329
2.8527 | -96.728
2.8565 | -96.7671
2.8602 | -96.7315
2.864 | -96.6255
2.8677 | -96.4735
2.8715 | -96.3209
2.8753 | -96.2054
2.879 | -96.1603
2.8828 | -96.1881
2.8865 | -96.2566
2.8903 | -96.3198
2.8941 | -96.3388
2.8978 | -96.3104
2.9016 | -96.2594
2.9053 | -96.229
2.9091 | -96.2391
2.9128 | -96.29
2.9166 | -96.3817
2.9204 | -96.4918
2.9241 | -96.6008
2.9279 | -96.6907
2.9316 | -96.7374
2.9354 | -96.7475
2.9392 | -96.7235
2.9429 | -96.6889
2.9467 | -96.6596
2.9504 | -96.6438
2.9542 | -96.6356
2.958 | -96.6417
2.9617 | -96.6386
2.9655 | -96.5896
2.9692 | -96.498
2.973 | -96.3742
2.9767 | -96.262
2.9805 | -96.204
2.9843 | -96.2096
2.988 | -96.2854
2.9918 | -96.4071
2.9955 | -96.552
2.9993 | -96.7162
3.0031 | -96.9064
3.0068 | -97.1082
3.0106 | -97.302
3.0143 | -97.4224
3.0181 | -97.3886
3.0218 | -97.2166
3.0256 | -96.9582
3.0294 | -96.7058
3.0331 | -96.5293
3.0369 | -96.4481
3.0406 | -96.4574
3.0444 | -96.5149
3.0482 | -96.5596
3.0519 | -96.5356
3.0557 | -96.4423
3.0594 | -96.3153
3.0632 | -96.2096
3.0669 | -96.1691
3.0707 | -96.2073
3.0745 | -96.3345
3.0782 | -96.5375
3.082 | -96.7885
3.0857 | -97.0439
3.0895 | -97.2657
3.0933 | -97.4373
3.097 | -97.5449
3.1008 | -97.585
3.1045 | -97.5489
3.1083 | -97.4378
3.1121 | -97.2912
3.1158 | -97.1451
3.1196 | -97.0379
3.1233 | -96.9693
3.1271 | -96.9541
3.1308 | -96.9845
3.1346 | -97.0457
3.1384 | -97.1164
3.1421 | -97.1887
3.1459 | -97.2676
3.1496 | -97.3541
3.1534 | -97.446
3.1572 | -97.5333
3.1609 | -97.5957
3.1647 | -97.6231
3.1684 | -97.6016
3.1722 | -97.5457
3.1759 | -97.4384
3.1797 | -97.2992
3.1835 | -97.1593
3.1872 | -97.0526
3.191 | -96.9921
3.1947 | -96.968
3.1985 | -96.9567
3.2023 | -96.9411
3.206 | -96.9177
3.2098 | -96.8917
3.2135 | -96.873
3.2173 | -96.8775
3.221 | -96.8863
3.2248 | -96.8996
3.2286 | -96.9271
3.2323 | -96.9631
3.2361 | -97.0128
3.2398 | -97.0881
3.2436 | -97.1895
3.2474 | -97.3075
3.2511 | -97.4066
3.2549 | -97.4984
3.2586 | -97.5606
3.2624 | -97.5802
3.2662 | -97.564
3.2699 | -97.5167
3.2737 | -97.4619
3.2774 | -97.4029
3.2812 | -97.3497
3.2849 | -97.315
3.2887 | -97.2925
3.2925 | -97.2744
3.2962 | -97.2527
3.3 | -97.2259
3.3037 | -97.1941
3.3075 | -97.1524
3.3113 | -97.1265
3.315 | -97.141
3.3188 | -97.2062
3.3225 | -97.3054
3.3263 | -97.4233
3.33 | -97.5326
3.3338 | -97.6098
3.3376 | -97.6356
3.3413 | -97.6146
3.3451 | -97.5518
3.3488 | -97.4594
3.3526 | -97.363
3.3564 | -97.2909
3.3601 | -97.2519
3.3639 | -97.2584
3.3676 | -97.2946
3.3714 | -97.3571
3.3751 | -97.4261
3.3789 | -97.4896
3.3827 | -97.5423
3.3864 | -97.5871
3.3902 | -97.6216
3.3939 | -97.6408
3.3977 | -97.6398
3.4015 | -97.6239
3.4052 | -97.6065
3.409 | -97.5995
3.4127 | -97.6076
3.4165 | -97.6471
3.4202 | -97.7053
3.424 | -97.7645
3.4278 | -97.7861
3.4315 | -97.7667
3.4353 | -97.7127
3.439 | -97.6262
3.4428 | -97.5355
3.4466 | -97.4707
3.4503 | -97.4564
3.4541 | -97.5016
3.4578 | -97.5934
3.4616 | -97.7025
3.4654 | -97.7967
3.4691 | -97.8591
3.4729 | -97.8637
3.4766 | -97.8189
3.4804 | -97.7456
3.4841 | -97.6743
3.4879 | -97.6223
3.4917 | -97.6106
3.4954 | -97.6399
3.4992 | -97.6919
3.5029 | -97.7532
3.5067 | -97.819
3.5105 | -97.8732
3.5142 | -97.9121
3.518 | -97.9435
3.5217 | -97.9712
3.5255 | -97.9855
3.5292 | -98.0005
3.533 | -98.0207
3.5368 | -98.0487
3.5405 | -98.0725
3.5443 | -98.0927
3.548 | -98.1169
3.5518 | -98.1537
3.5556 | -98.2083
3.5593 | -98.2799
3.5631 | -98.3661
3.5668 | -98.4484
3.5706 | -98.5079
3.5743 | -98.536
3.5781 | -98.5178
3.5819 | -98.4554
3.5856 | -98.361
3.5894 | -98.2504
3.5931 | -98.141
3.5969 | -98.053
3.6007 | -97.9921
3.6044 | -97.9387
3.6082 | -97.9051
3.6119 | -97.8865
3.6157 | -97.8756
3.6195 | -97.8698
3.6232 | -97.8709
3.627 | -97.8584
3.6307 | -97.8269
3.6345 | -97.7854
3.6382 | -97.7469
3.642 | -97.7218
3.6458 | -97.7304
3.6495 | -97.7831
3.6533 | -97.8595
3.657 | -97.9343
3.6608 | -97.9985
3.6646 | -98.0523
3.6683 | -98.1097
3.6721 | -98.1697
3.6758 | -98.2086
3.6796 | -98.2099
3.6833 | -98.1631
3.6871 | -98.0633
3.6909 | -97.95
3.6946 | -97.8649
3.6984 | -97.8195
3.7021 | -97.8225
3.7059 | -97.8622
3.7097 | -97.9255
3.7134 | -98.0069
3.7172 | -98.0905
3.7209 | -98.1567
3.7247 | -98.1912
3.7284 | -98.2098
3.7322 | -98.2183
3.736 | -98.2109
3.7397 | -98.1973
3.7435 | -98.1863
3.7472 | -98.18
3.751 | -98.1753
3.7548 | -98.1816
3.7585 | -98.2043
3.7623 | -98.2582
3.766 | -98.3491
3.7698 | -98.4602
3.7735 | -98.5864
3.7773 | -98.7106
3.7811 | -98.8192
3.7848 | -98.871
3.7886 | -98.8493
3.7923 | -98.6959
3.7961 | -98.5903
3.7999 | -98.4704
3.8036 | -98.3586
3.8074 | -98.2715
3.8111 | -98.2235
3.8149 | -98.2233
3.8187 | -98.2528
3.8224 | -98.3177
3.8262 | -98.4084
3.8299 | -98.5231
3.8337 | -98.6594
3.8374 | -98.7944
3.8412 | -98.9094
3.845 | -99.0031
3.8487 | -99.0762
3.8525 | -99.1223
3.8562 | -99.1476
3.86 | -99.1567
3.8638 | -99.1437
3.8675 | -99.1215
3.8713 | -99.0834
3.875 | -99.0353
3.8788 | -98.9965
3.8825 | -98.9687
3.8863 | -98.9432
3.8901 | -98.9319
3.8938 | -98.9211
3.8976 | -98.8952
3.9013 | -98.8517
3.9051 | -98.7898
3.9089 | -98.717
3.9126 | -98.645
3.9164 | -98.5858
3.9201 | -98.545
3.9239 | -98.5274
3.9276 | -98.5278
3.9314 | -98.5391
3.9352 | -98.5594
3.9389 | -98.5824
3.9427 | -98.5889
3.9464 | -98.5749
3.9502 | -98.5407
3.954 | -98.4993
3.9577 | -98.4662
3.9615 | -98.4612
3.9652 | -98.4973
3.969 | -98.5724
3.9728 | -98.6767
3.9765 | -98.7902
3.9803 | -98.8969
3.984 | -98.979
3.9878 | -99.0241
3.9915 | -99.0276
3.9953 | -98.9926
3.9991 | -98.9422
4.0028 | -98.8838
4.0066 | -98.8335
4.0103 | -98.8061
4.0141 | -98.8079
4.0179 | -98.8381
4.0216 | -98.8976
4.0254 | -98.9731
4.0291 | -99.0531
4.0329 | -99.1358
4.0366 | -99.2155
4.0404 | -99.2855
4.0442 | -99.3447
4.0479 | -99.4022
4.0517 | -99.4573
4.0554 | -99.5037
4.0592 | -99.532
4.063 | -99.5499
4.0667 | -99.5682
4.0705 | -99.5829
4.0742 | -99.5941
4.078 | -99.5906
4.0817 | -99.5841
4.0855 | -99.5706
4.0893 | -99.5508
4.093 | -99.5221
4.0968 | -99.4799
4.1005 | -99.4251
4.1043 | -99.3744
4.1081 | -99.3214
4.1118 | -99.2733
4.1156 | -99.2412
4.1193 | -99.231
4.1231 | -99.25
4.1269 | -99.2788
4.1306 | -99.3218
4.1344 | -99.3675
4.1381 | -99.4003
4.1419 | -99.4215
4.1456 | -99.4285
4.1494 | -99.4318
4.1532 | -99.4433
4.1569 | -99.4699
4.1607 | -99.5227
4.1644 | -99.5908
4.1682 | -99.6644
4.172 | -99.7401
4.1757 | -99.7903
4.1795 | -99.8212
4.1832 | -99.8234
4.187 | -99.7836
4.1907 | -99.7206
4.1945 | -99.6368
4.1983 | -99.5511
4.202 | -99.4774
4.2058 | -99.4152
4.2095 | -99.3767
4.2133 | -99.3783
4.2171 | -99.4078
4.2208 | -99.4608
4.2246 | -99.5277
4.2283 | -99.6041
4.2321 | -99.6805
4.2358 | -99.7524
4.2396 | -99.8363
4.2434 | -99.9249
4.2471 | -100.0225
4.2509 | -100.1166
4.2546 | -100.1815
4.2584 | -100.2085
4.2622 | -100.1971
4.2659 | -100.1501
4.2697 | -100.0809
4.2734 | -99.9991
4.2772 | -99.9149
4.2809 | -99.8374
4.2847 | -99.7698
4.2885 | -99.7133
4.2922 | -99.6796
4.296 | -99.6666
4.2997 | -99.6634
4.3035 | -99.6663
4.3073 | -99.6669
4.311 | -99.658
4.3148 | -99.6554
4.3185 | -99.6599
4.3223 | -99.6653
4.3261 | -99.6743
4.3298 | -99.6861
4.3336 | -99.7066
4.3373 | -99.7336
4.3411 | -99.7588
4.3448 | -99.7765
4.3486 | -99.7911
4.3524 | -99.8018
4.3561 | -99.81
4.3599 | -99.807
4.3636 | -99.8088
4.3674 | -99.8177
4.3712 | -99.8265
4.3749 | -99.8412
4.3787 | -99.8584
4.3824 | -99.8637
4.3862 | -99.8716
4.3899 | -99.8784
4.3937 | -99.8772
4.3975 | -99.8745
4.4012 | -99.865
4.405 | -99.845
4.4087 | -99.8229
4.4125 | -99.798
4.4163 | -99.7797
4.42 | -99.7645
4.4238 | -99.7529
4.4275 | -99.7469
4.4313 | -99.737
4.435 | -99.7151
4.4388 | -99.6928
4.4426 | -99.6642
4.4463 | -99.6391
4.4501 | -99.6235
4.4538 | -99.6194
4.4576 | -99.6179
4.4614 | -99.6347
4.4651 | -99.664
4.4689 | -99.6863
4.4726 | -99.6982
4.4764 | -99.6965
4.4802 | -99.6838
4.4839 | -99.6777
4.4877 | -99.6758
4.4914 | -99.6937
4.4952 | -99.7229
4.4989 | -99.7518
4.5027 | -99.7858
4.5065 | -99.8103
4.5102 | -99.8178
4.514 | -99.8138
4.5177 | -99.8009
4.5215 | -99.7987
4.5253 | -99.8244
4.529 | -99.8786
4.5328 | -99.9539
4.5365 | -100.0463
4.5403 | -100.1493
4.544 | -100.2526
4.5478 | -100.3451
4.5516 | -100.4118
4.5553 | -100.448
4.5591 | -100.4701
4.5628 | -100.4694
4.5666 | -100.4558
4.5704 | -100.4355
4.5741 | -100.4076
4.5779 | -100.3792
4.5816 | -100.3439
4.5854 | -100.2955
4.5891 | -100.2393
4.5929 | -100.1788
4.5967 | -100.1152
4.6004 | -100.0447
4.6042 | -99.9691
4.6079 | -99.8914
4.6117 | -99.8219
4.6155 | -99.7587
4.6192 | -99.7154
4.623 | -99.706
4.6267 | -99.7272
4.6305 | -99.7695
4.6342 | -99.8285
4.638 | -99.8978
4.6418 | -99.9705
4.6455 | -100.0421
4.6493 | -100.1142
4.653 | -100.2051
4.6568 | -100.3084
4.6606 | -100.4113
4.6643 | -100.5023
4.6681 | -100.565
4.6718 | -100.608
4.6756 | -100.6236
4.6794 | -100.6287
4.6831 | -100.6323
4.6869 | -100.6457
4.6906 | -100.6703
4.6944 | -100.6988
4.6981 | -100.7297
4.7019 | -100.7638
4.7057 | -100.7939
4.7094 | -100.8133
4.7132 | -100.8282
4.7169 | -100.8279
4.7207 | -100.8184
4.7245 | -100.8004
4.7282 | -100.7828
4.732 | -100.7602
4.7357 | -100.737
4.7395 | -100.7269
4.7432 | -100.7178
4.747 | -100.7126
4.7508 | -100.7104
4.7545 | -100.7178
4.7583 | -100.7378
4.762 | -100.7657
4.7658 | -100.8068
4.7696 | -100.8623
4.7733 | -100.9248
4.7771 | -100.9851
4.7808 | -101.0403
4.7846 | -101.0791
4.7883 | -101.0979
4.7921 | -101.0852
4.7959 | -101.0338
4.7996 | -100.9501
4.8034 | -100.8507
4.8071 | -100.7426
4.8109 | -100.6355
4.8147 | -100.5466
4.8184 | -100.4883
4.8222 | -100.4607
4.8259 | -100.4682
4.8297 | -100.5105
4.8335 | -100.5758
4.8372 | -100.635
4.841 | -100.6724
4.8447 | -100.6727
4.8485 | -100.6425
4.8522 | -100.5843
4.856 | -100.5146
4.8598 | -100.4474
4.8635 | -100.3907
4.8673 | -100.3516
4.871 | -100.3353
4.8748 | -100.3306
4.8786 | -100.3593
4.8823 | -100.3785
4.8861 | -100.4124
4.8898 | -100.453
4.8936 | -100.4929
4.8973 | -100.5365
4.9011 | -100.5776
4.9049 | -100.6153
4.9086 | -100.6513
4.9124 | -100.6884
4.9161 | -100.7161
4.9199 | -100.7303
4.9237 | -100.731
4.9274 | -100.7196
4.9312 | -100.6966
4.9349 | -100.6689
4.9387 | -100.638
4.9424 | -100.6109
4.9462 | -100.59
4.95 | -100.582
4.9537 | -100.5835
4.9575 | -100.5798
4.9612 | -100.5674
4.965 | -100.5483
4.9688 | -100.52
4.9725 | -100.4831
4.9763 | -100.437
4.98 | -100.3839
4.9838 | -100.3229
4.9875 | -100.2633
4.9913 | -100.2079
4.9951 | -100.1614
4.9988 | -100.1266
5.0026 | -100.0996
5.0063 | -100.0855
5.0101 | -100.0895
5.0139 | -100.1036
5.0176 | -100.1245
5.0214 | -100.1506
5.0251 | -100.1699
5.0289 | -100.1771
5.0327 | -100.1684
5.0364 | -100.1455
5.0402 | -100.1041
5.0439 | -100.0521
5.0477 | -100.0051
5.0514 | -99.9685
5.0552 | -99.9448
5.059 | -99.9507
5.0627 | -99.9924
5.0665 | -100.0636
5.0702 | -100.1626
5.074 | -100.284
5.0778 | -100.4188
5.0815 | -100.5554
5.0853 | -100.6915
5.089 | -100.8185
5.0928 | -100.9377
5.0965 | -101.0434
5.1003 | -101.125
5.1041 | -101.1868
5.1078 | -101.2356
5.1116 | -101.2678
5.1153 | -101.2865
5.1191 | -101.2941
5.1229 | -101.2866
5.1266 | -101.2624
5.1304 | -101.2385
5.1341 | -101.2189
5.1379 | -101.1996
5.1416 | -101.1828
5.1454 | -101.1784
5.1492 | -101.189
5.1529 | -101.2019
5.1567 | -101.2223
5.1604 | -101.2442
5.1642 | -101.2581
5.168 | -101.2594
5.1717 | -101.2437
5.1755 | -101.2077
5.1792 | -101.1599
5.183 | -101.1119
5.1868 | -101.0605
5.1905 | -101.0162
5.1943 | -100.9896
5.198 | -100.9792
5.2018 | -100.99
5.2055 | -101.0247
5.2093 | -101.0756
5.2131 | -101.1469
5.2168 | -101.2261
5.2206 | -101.3082
5.2243 | -101.3828
5.2281 | -101.4409
5.2319 | -101.4804
5.2356 | -101.4838
5.2394 | -101.4548
5.2431 | -101.4042
5.2469 | -101.3457
5.2506 | -101.2909
5.2544 | -101.2479
5.2582 | -101.2169
5.2619 | -101.1959
5.2657 | -101.1901
5.2694 | -101.1968
5.2732 | -101.201
5.277 | -101.2073
5.2807 | -101.2057
5.2845 | -101.192
5.2882 | -101.1687
5.292 | -101.1454
5.2957 | -101.1169
5.2995 | -101.0816
5.3033 | -101.0432
5.307 | -100.9959
5.3108 | -100.9466
5.3145 | -100.9088
5.3183 | -100.8865
5.3221 | -100.8761
5.3258 | -100.885
5.3296 | -100.9141
5.3333 | -100.9696
5.3371 | -101.0523
5.3408 | -101.1455
5.3446 | -101.2413
5.3484 | -101.3357
5.3521 | -101.4266
5.3559 | -101.499
5.3596 | -101.5627
5.3634 | -101.6146
5.3672 | -101.6561
5.3709 | -101.6843
5.3747 | -101.7101
5.3784 | -101.7197
5.3822 | -101.7105
5.386 | -101.6875
5.3897 | -101.6458
5.3935 | -101.5826
5.3972 | -101.5096
5.401 | -101.4433
5.4047 | -101.3864
5.4085 | -101.3428
5.4123 | -101.3079
5.416 | -101.2829
5.4198 | -101.2871
5.4235 | -101.2789
5.4273 | -101.2737
5.4311 | -101.2638
5.4348 | -101.2467
5.4386 | -101.2278
5.4423 | -101.1906
5.4461 | -101.1534
5.4498 | -101.1158
5.4536 | -101.0743
5.4574 | -101.037
5.4611 | -101.016
5.4649 | -101.014
5.4686 | -101.0328
5.4724 | -101.0778
5.4762 | -101.1483
5.4799 | -101.2361
5.4837 | -101.3365
5.4874 | -101.442
5.4912 | -101.5468
5.4949 | -101.6359
5.4987 | -101.7121
5.5025 | -101.7686
5.5062 | -101.7971
5.51 | -101.8012
5.5137 | -101.7836
5.5175 | -101.7428
5.5213 | -101.6929
5.525 | -101.6367
5.5288 | -101.578
5.5325 | -101.507
5.5363 | -101.4317
5.5401 | -101.3621
5.5438 | -101.304
5.5476 | -101.2563
5.5513 | -101.222
5.5551 | -101.2012
5.5588 | -101.1922
5.5626 | -101.1975
5.5664 | -101.2158
5.5701 | -101.2366
5.5739 | -101.2588
5.5776 | -101.2893
5.5814 | -101.3183
5.5852 | -101.3372
5.5889 | -101.3551
5.5927 | -101.3777
5.5964 | -101.3965
5.6002 | -101.4223
5.6039 | -101.4559
5.6077 | -101.4846
5.6115 | -101.51
5.6152 | -101.5417
5.619 | -101.5733
5.6227 | -101.6115
5.6265 | -101.6429
5.6303 | -101.6572
5.634 | -101.6598
5.6378 | -101.6533
5.6415 | -101.6365
5.6453 | -101.6113
5.649 | -101.5767
5.6528 | -101.5401
5.6566 | -101.5096
5.6603 | -101.4855
5.6641 | -101.4625
5.6678 | -101.4535
5.6716 | -101.4533
5.6754 | -101.4533
5.6791 | -101.4626
5.6829 | -101.4822
5.6866 | -101.5015
5.6904 | -101.5213
5.6942 | -101.5427
5.6979 | -101.563
5.7017 | -101.5733
5.7054 | -101.5744
5.7092 | -101.5514
5.7129 | -101.5269
5.7167 | -101.4978
5.7205 | -101.4619
5.7242 | -101.4225
5.728 | -101.3921
5.7317 | -101.3637
5.7355 | -101.3451
5.7393 | -101.3323
5.743 | -101.329
5.7468 | -101.3445
5.7505 | -101.3719
5.7543 | -101.4146
5.758 | -101.4758
5.7618 | -101.5497
5.7656 | -101.6316
5.7693 | -101.726
5.7731 | -101.8308
5.7768 | -101.9376
5.7806 | -102.0402
5.7844 | -102.1297
5.7881 | -102.2043
5.7919 | -102.2571
5.7956 | -102.2929
5.7994 | -102.3021
5.8031 | -102.2905
5.8069 | -102.2551
5.8107 | -102.2062
5.8144 | -102.1496
5.8182 | -102.0912
5.8219 | -102.0401
5.8257 | -102.0056
5.8295 | -101.9857
5.8332 | -101.9765
5.837 | -101.9841
5.8407 | -101.9998
5.8445 | -102.0109
5.8483 | -102.0212
5.852 | -102.0151
5.8558 | -101.9904
5.8595 | -101.9571
5.8633 | -101.9184
5.867 | -101.8706
5.8708 | -101.8257
5.8746 | -101.7883
5.8783 | -101.756
5.8821 | -101.7325
5.8858 | -101.7166
5.8896 | -101.709
5.8934 | -101.7059
5.8971 | -101.7115
5.9009 | -101.7147
5.9046 | -101.7208
5.9084 | -101.718
5.9121 | -101.7077
5.9159 | -101.6943
5.9197 | -101.6744
5.9234 | -101.6564
5.9272 | -101.6469
5.9309 | -101.6456
5.9347 | -101.657
5.9385 | -101.6841
5.9422 | -101.7233
5.946 | -101.7697
5.9497 | -101.8201
5.9535 | -101.8793
5.9572 | -101.9406
5.961 | -101.977
5.9648 | -102.0139
5.9685 | -102.0429
5.9723 | -102.0543
5.976 | -102.051
5.9798 | -102.0337
5.9836 | -102.0045
5.9873 | -101.9603
5.9911 | -101.9111
5.9948 | -101.8545
5.9986 | -101.8004
6.0023 | -101.7514
6.0061 | -101.7065
6.0099 | -101.6674
6.0136 | -101.6308
6.0174 | -101.598
6.0211 | -101.5747
6.0249 | -101.5569
6.0287 | -101.5462
6.0324 | -101.5399
6.0362 | -101.5361
6.0399 | -101.5363
6.0437 | -101.5394
6.0475 | -101.5467
6.0512 | -101.5605
6.055 | -101.5856
6.0587 | -101.6286
6.0625 | -101.6857
6.0662 | -101.7596
6.07 | -101.8473
6.0738 | -101.9481
6.0775 | -102.0563
6.0813 | -102.1668
6.085 | -102.284
6.0888 | -102.3983
6.0926 | -102.4924
6.0963 | -102.566
6.1001 | -102.6176
6.1038 | -102.6525
6.1076 | -102.6626
6.1113 | -102.6428
6.1151 | -102.598
6.1189 | -102.534
6.1226 | -102.4537
6.1264 | -102.3713
6.1301 | -102.2895
6.1339 | -102.2104
6.1377 | -102.1384
6.1414 | -102.0791
6.1452 | -102.0363
6.1489 | -101.9981
6.1527 | -101.9664
6.1564 | -101.9523
6.1602 | -101.9476
6.164 | -101.9509
6.1677 | -101.9597
6.1715 | -101.9745
6.1752 | -101.9919
6.179 | -102.0157
6.1828 | -102.0525
6.1865 | -102.1022
6.1903 | -102.1635
6.194 | -102.2285
6.1978 | -102.2969
6.2016 | -102.367
6.2053 | -102.4398
6.2091 | -102.5092
6.2128 | -102.5611
6.2166 | -102.5994
6.2203 | -102.6178
6.2241 | -102.6178
6.2279 | -102.6012
6.2316 | -102.5703
6.2354 | -102.5265
6.2391 | -102.4792
6.2429 | -102.4402
6.2467 | -102.4023
6.2504 | -102.3659
6.2542 | -102.3416
6.2579 | -102.3296
6.2617 | -102.3214
6.2654 | -102.3176
6.2692 | -102.3156
6.273 | -102.3103
6.2767 | -102.3024
6.2805 | -102.2991
6.2842 | -102.2904
6.288 | -102.2829
6.2918 | -102.2802
6.2955 | -102.2835
6.2993 | -102.2916
6.303 | -102.3025
6.3068 | -102.3279
6.3105 | -102.3624
6.3143 | -102.3983
6.3181 | -102.4365
6.3218 | -102.4771
6.3256 | -102.5201
6.3293 | -102.5621
6.3331 | -102.5971
6.3369 | -102.6204
6.3406 | -102.6367
6.3444 | -102.6488
6.3481 | -102.6565
6.3519 | -102.6602
6.3556 | -102.6652
6.3594 | -102.6688
6.3632 | -102.6709
6.3669 | -102.673
6.3707 | -102.6741
6.3744 | -102.6729
6.3782 | -102.6739
6.382 | -102.6763
6.3857 | -102.6718
6.3895 | -102.6636
6.3932 | -102.6579
6.397 | -102.655
6.4008 | -102.6536
6.4045 | -102.6542
6.4083 | -102.6618
6.412 | -102.6758
6.4158 | -102.6903
6.4195 | -102.7016
6.4233 | -102.7156
6.4271 | -102.7266
6.4308 | -102.7351
6.4346 | -102.7342
6.4383 | -102.7243
6.4421 | -102.7088
6.4459 | -102.6864
6.4496 | -102.6656
6.4534 | -102.6399
6.4571 | -102.608
6.4609 | -102.5769
6.4646 | -102.5428
6.4684 | -102.5156
6.4722 | -102.4979
6.4759 | -102.4817
6.4797 | -102.4714
6.4834 | -102.469
6.4872 | -102.4714
6.491 | -102.4837
6.4947 | -102.4979
6.4985 | -102.5183
6.5022 | -102.5444
6.506 | -102.5628
6.5097 | -102.5867
6.5135 | -102.6074
6.5173 | -102.6232
6.521 | -102.6373
6.5248 | -102.6531
6.5285 | -102.6639
6.5323 | -102.676
6.5361 | -102.6875
6.5398 | -102.6958
6.5436 | -102.7057
6.5473 | -102.7091
6.5511 | -102.7134
6.5549 | -102.7191
6.5586 | -102.7191
6.5624 | -102.7205
6.5661 | -102.7135
6.5699 | -102.706
6.5736 | -102.7021
6.5774 | -102.6978
6.5812 | -102.6997
6.5849 | -102.7067
6.5887 | -102.7159
6.5924 | -102.723
6.5962 | -102.7321
6.6 | -102.7423
6.6037 | -102.7522
6.6075 | -102.7512
6.6112 | -102.7389
6.615 | -102.7193
6.6187 | -102.6912
6.6225 | -102.6542
6.6263 | -102.6126
6.63 | -102.565
6.6338 | -102.5134
6.6375 | -102.4578
6.6413 | -102.4039
6.6451 | -102.3527
6.6488 | -102.3042
6.6526 | -102.26
6.6563 | -102.2243
6.6601 | -102.1941
6.6638 | -102.1672
6.6676 | -102.154
6.6714 | -102.1498
6.6751 | -102.1569
6.6789 | -102.1755
6.6826 | -102.1993
6.6864 | -102.2321
6.6902 | -102.2826
6.6939 | -102.3462
6.6977 | -102.4111
6.7014 | -102.4828
6.7052 | -102.557
6.7089 | -102.6321
6.7127 | -102.6994
6.7165 | -102.7566
6.7202 | -102.7996
6.724 | -102.8277
6.7277 | -102.8431
6.7315 | -102.8543
6.7353 | -102.8588
6.739 | -102.858
6.7428 | -102.852
6.7465 | -102.8489
6.7503 | -102.8511
6.7541 | -102.8457
6.7578 | -102.837
6.7616 | -102.8343
6.7653 | -102.8355
6.7691 | -102.8306
6.7728 | -102.8224
6.7766 | -102.8112
6.7804 | -102.7927
6.7841 | -102.7686
6.7879 | -102.7395
6.7916 | -102.6971
6.7954 | -102.6603
6.7992 | -102.6341
6.8029 | -102.6132
6.8067 | -102.6001
6.8104 | -102.5913
6.8142 | -102.5927
6.8179 | -102.6071
6.8217 | -102.6241
6.8255 | -102.6456
6.8292 | -102.6737
6.833 | -102.6998
6.8367 | -102.7195
6.8405 | -102.7412
6.8443 | -102.7653
6.848 | -102.7858
6.8518 | -102.8038
6.8555 | -102.8255
6.8593 | -102.8421
6.863 | -102.8497
6.8668 | -102.852
6.8706 | -102.8518
6.8743 | -102.851
6.8781 | -102.8492
6.8818 | -102.8427
6.8856 | -102.8324
6.8894 | -102.8188
6.8931 | -102.8076
6.8969 | -102.7876
6.9006 | -102.7582
6.9044 | -102.728
6.9082 | -102.7003
6.9119 | -102.6764
6.9157 | -102.6481
6.9194 | -102.6205
6.9232 | -102.6055
6.9269 | -102.5986
6.9307 | -102.6012
6.9345 | -102.6175
6.9382 | -102.6514
6.942 | -102.6938
6.9457 | -102.7388
6.9495 | -102.7817
6.9533 | -102.8349
6.957 | -102.8932
6.9608 | -102.9474
6.9645 | -103.0029
6.9683 | -103.0569
6.972 | -103.1003
6.9758 | -103.1327
6.9796 | -103.1617
6.9833 | -103.1843
6.9871 | -103.195
6.9908 | -103.2009
6.9946 | -103.1947
6.9984 | -103.1836
7.0021 | -103.1682
7.0059 | -103.1534
7.0096 | -103.1397
7.0134 | -103.1229
7.0171 | -103.1039
7.0209 | -103.0782
7.0247 | -103.0431
7.0284 | -103.0037
7.0322 | -102.965
7.0359 | -102.9329
7.0397 | -102.9048
7.0435 | -102.883
7.0472 | -102.8896
7.051 | -102.887
7.0547 | -102.8923
7.0585 | -102.8991
7.0622 | -102.9152
7.066 | -102.9367
7.0698 | -102.9555
7.0735 | -102.9795
7.0773 | -103.0106
7.081 | -103.0454
7.0848 | -103.0738
7.0886 | -103.1001
7.0923 | -103.1161
7.0961 | -103.1211
7.0998 | -103.1187
7.1036 | -103.1031
7.1074 | -103.0788
7.1111 | -103.0392
7.1149 | -102.9983
7.1186 | -102.9583
7.1224 | -102.9243
7.1261 | -102.8968
7.1299 | -102.8741
7.1337 | -102.8548
7.1374 | -102.8397
7.1412 | -102.8261
7.1449 | -102.8194
7.1487 | -102.8225
7.1525 | -102.8309
7.1562 | -102.8435
7.16 | -102.8631
7.1637 | -102.8875
7.1675 | -102.9094
7.1712 | -102.9276
7.175 | -102.952
7.1788 | -102.9764
7.1825 | -102.9954
7.1863 | -103.0105
7.19 | -103.0211
7.1938 | -103.0341
7.1976 | -103.0473
7.2013 | -103.0611
7.2051 | -103.0678
7.2088 | -103.0729
7.2126 | -103.0736
7.2164 | -103.0769
7.2201 | -103.0883
7.2239 | -103.0955
7.2276 | -103.1055
7.2314 | -103.1187
7.2351 | -103.1366
7.2389 | -103.1587
7.2427 | -103.183
7.2464 | -103.2071
7.2502 | -103.2346
7.2539 | -103.2566
7.2577 | -103.2727
7.2615 | -103.2886
7.2652 | -103.3057
7.269 | -103.3174
7.2727 | -103.3245
7.2765 | -103.3302
7.2802 | -103.3341
7.284 | -103.3324
7.2878 | -103.3323
7.2915 | -103.3294
7.2953 | -103.3218
7.299 | -103.3106
7.3028 | -103.3001
7.3066 | -103.2898
7.3103 | -103.2836
7.3141 | -103.2827
7.3178 | -103.2846
7.3216 | -103.2903
7.3253 | -103.2933
7.3291 | -103.2947
7.3329 | -103.2969
7.3366 | -103.3069
7.3404 | -103.3137
7.3441 | -103.3145
7.3479 | -103.3156
7.3517 | -103.3162
7.3554 | -103.3156
7.3592 | -103.3129
7.3629 | -103.3153
7.3667 | -103.318
7.3704 | -103.3237
7.3742 | -103.3291
7.378 | -103.3446
7.3817 | -103.3668
7.3855 | -103.392
7.3892 | -103.422
7.393 | -103.4589
7.3968 | -103.4974
7.4005 | -103.5317
7.4043 | -103.5672
7.408 | -103.6002
7.4118 | -103.6211
7.4156 | -103.6381
7.4193 | -103.6484
7.4231 | -103.6487
7.4268 | -103.6336
7.4306 | -103.617
7.4343 | -103.595
7.4381 | -103.5669
7.4419 | -103.5377
7.4456 | -103.5009
7.4494 | -103.4714
7.4531 | -103.4477
7.4569 | -103.4223
7.4607 | -103.3996
7.4644 | -103.3813
7.4682 | -103.3728
7.4719 | -103.3701
7.4757 | -103.3736
7.4794 | -103.3817
7.4832 | -103.395
7.487 | -103.4133
7.4907 | -103.432
7.4945 | -103.4491
7.4982 | -103.4652
7.502 | -103.4786
7.5058 | -103.4899
7.5095 | -103.5024
7.5133 | -103.5064
7.517 | -103.5043
7.5208 | -103.5018
7.5245 | -103.4925
7.5283 | -103.4785
7.5321 | -103.4639
7.5358 | -103.4553
7.5396 | -103.4389
7.5433 | -103.4212
7.5471 | -103.4048
7.5509 | -103.3935
7.5546 | -103.3801
7.5584 | -103.3713
7.5621 | -103.3668
7.5659 | -103.3656
7.5697 | -103.3718
7.5734 | -103.3849
7.5772 | -103.3957
7.5809 | -103.4077
7.5847 | -103.4191
7.5884 | -103.4069
7.5922 | -103.4036
7.596 | -103.3894
7.5997 | -103.3693
7.6035 | -103.3497
7.6072 | -103.3305
7.611 | -103.3091
7.6148 | -103.287
7.6185 | -103.2684
7.6223 | -103.2525
7.626 | -103.2429
7.6298 | -103.2381
7.6335 | -103.2401
7.6373 | -103.2492
7.6411 | -103.2662
7.6448 | -103.2872
7.6486 | -103.3109
7.6523 | -103.34
7.6561 | -103.371
7.6599 | -103.4073
7.6636 | -103.4411
7.6674 | -103.4798
7.6711 | -103.5139
7.6749 | -103.5502
7.6786 | -103.5858
7.6824 | -103.603
7.6862 | -103.6016
7.6899 | -103.6043
7.6937 | -103.6075
7.6974 | -103.6043
7.7012 | -103.5935
7.705 | -103.5718
7.7087 | -103.5407
7.7125 | -103.5111
7.7162 | -103.4722
7.72 | -103.4281
7.7237 | -103.3889
7.7275 | -103.3542
7.7313 | -103.329
7.735 | -103.3121
7.7388 | -103.3056
7.7425 | -103.3105
7.7463 | -103.3212
7.7501 | -103.3388
7.7538 | -103.3701
7.7576 | -103.4064
7.7613 | -103.4505
7.7651 | -103.4925
7.7689 | -103.5311
7.7726 | -103.5674
7.7764 | -103.5978
7.7801 | -103.6266
7.7839 | -103.6489
7.7876 | -103.6721
7.7914 | -103.6915
7.7952 | -103.7043
7.7989 | -103.7212
7.8027 | -103.7335
7.8064 | -103.7274
7.8102 | -103.7128
7.814 | -103.6983
7.8177 | -103.6854
7.8215 | -103.6797
7.8252 | -103.6777
7.829 | -103.6759
7.8327 | -103.6858
7.8365 | -103.7054
7.8403 | -103.7228
7.844 | -103.7409
7.8478 | -103.7679
7.8515 | -103.8024
7.8553 | -103.8329
7.8591 | -103.8611
7.8628 | -103.8911
7.8666 | -103.9099
7.8703 | -103.9136
7.8741 | -103.9066
7.8778 | -103.8884
7.8816 | -103.8691
7.8854 | -103.8448
7.8891 | -103.8141
7.8929 | -103.7762
7.8966 | -103.7308
7.9004 | -103.6848
7.9042 | -103.6392
7.9079 | -103.598
7.9117 | -103.5679
7.9154 | -103.5411
7.9192 | -103.529
7.923 | -103.5268
7.9267 | -103.5326
7.9305 | -103.5451
7.9342 | -103.5619
7.938 | -103.5838
7.9417 | -103.6065
7.9455 | -103.6288
7.9493 | -103.6522
7.953 | -103.6716
7.9568 | -103.6914
7.9605 | -103.7147
7.9643 | -103.7249
7.9681 | -103.725
7.9718 | -103.728
7.9756 | -103.7258
7.9793 | -103.7171
7.9831 | -103.7063
7.9868 | -103.6955
7.9906 | -103.6851
7.9944 | -103.6709
7.9981 | -103.6539
8.0019 | -103.6367
8.0056 | -103.6198
8.0094 | -103.606
8.0132 | -103.5966
8.0169 | -103.5855
8.0207 | -103.573
8.0244 | -103.5597
8.0282 | -103.55
8.0319 | -103.5414
8.0357 | -103.5274
8.0395 | -103.5128
8.0432 | -103.4975
8.047 | -103.4847
8.0507 | -103.4738
8.0545 | -103.4596
8.0583 | -103.4486
8.062 | -103.4385
8.0658 | -103.4292
8.0695 | -103.427
8.0733 | -103.4329
8.077 | -103.4433
8.0808 | -103.4603
8.0846 | -103.4863
8.0883 | -103.5174
8.0921 | -103.5584
8.0958 | -103.6084
8.0996 | -103.6719
8.1034 | -103.7382
8.1071 | -103.8053
8.1109 | -103.8798
8.1146 | -103.9518
8.1184 | -104.0189
8.1222 | -104.078
8.1259 | -104.1312
8.1297 | -104.1703
8.1334 | -104.1515
8.1372 | -104.1605
8.1409 | -104.1577
8.1447 | -104.1432
8.1485 | -104.1111
8.1522 | -104.0732
8.156 | -104.0264
8.1597 | -103.9764
8.1635 | -103.925
8.1673 | -103.8721
8.171 | -103.8237
8.1748 | -103.7881
8.1785 | -103.758
8.1823 | -103.7392
8.186 | -103.7312
8.1898 | -103.7322
8.1936 | -103.7429
8.1973 | -103.7648
8.2011 | -103.8024
8.2048 | -103.85
8.2086 | -103.9025
8.2124 | -103.9531
8.2161 | -104.0079
8.2199 | -104.0631
8.2236 | -104.1169
8.2274 | -104.1674
8.2311 | -104.212
8.2349 | -104.2519
8.2387 | -104.2868
8.2424 | -104.314
8.2462 | -104.3322
8.2499 | -104.3462
8.2537 | -104.3491
8.2575 | -104.3429
8.2612 | -104.3308
8.265 | -104.3101
8.2687 | -104.284
8.2725 | -104.2593
8.2763 | -104.2307
8.28 | -104.2
8.2838 | -104.1669
8.2875 | -104.1343
8.2913 | -104.0965
8.295 | -104.0578
8.2988 | -104.0208
8.3026 | -103.9884
8.3063 | -103.9524
8.3101 | -103.9016
8.3138 | -103.8608
8.3176 | -103.8256
8.3214 | -103.7873
8.3251 | -103.7471
8.3289 | -103.7048
8.3326 | -103.6651
8.3364 | -103.6247
8.3401 | -103.5803
8.3439 | -103.5363
8.3477 | -103.4967
8.3514 | -103.4616
8.3552 | -103.4278
8.3589 | -103.403
8.3627 | -103.391
8.3665 | -103.3915
8.3702 | -103.4027
8.374 | -103.4265
8.3777 | -103.4665
8.3815 | -103.5126
8.3852 | -103.5678
8.389 | -103.6311
8.3928 | -103.6989
8.3965 | -103.7668
8.4003 | -103.8351
8.404 | -103.9066
8.4078 | -103.9678
8.4116 | -104.021
8.4153 | -104.0748
8.4191 | -104.118
8.4228 | -104.1493
8.4266 | -104.175
8.4304 | -104.1985
8.4341 | -104.221
8.4379 | -104.2246
8.4416 | -104.2199
8.4454 | -104.2171
8.4491 | -104.212
8.4529 | -104.2083
8.4567 | -104.2036
8.4604 | -104.2041
8.4642 | -104.2058
8.4679 | -104.2074
8.4717 | -104.2105
8.4755 | -104.2179
8.4792 | -104.2307
8.483 | -104.2408
8.4867 | -104.2482
8.4905 | -104.2582
8.4942 | -104.2724
8.498 | -104.2851
8.5018 | -104.2942
8.5055 | -104.3054
8.5093 | -104.3068
8.513 | -104.304
8.5168 | -104.2996
8.5206 | -104.2925
8.5243 | -104.2817
8.5281 | -104.27
8.5318 | -104.266
8.5356 | -104.2608
8.5393 | -104.2562
8.5431 | -104.2564
8.5469 | -104.2628
8.5506 | -104.2775
8.5544 | -104.3011
8.5581 | -104.3349
8.5619 | -104.3714
8.5657 | -104.4018
8.5694 | -104.4326
8.5732 | -104.4678
8.5769 | -104.4987
8.5807 | -104.5232
8.5844 | -104.5406
8.5882 | -104.5502
8.592 | -104.5479
8.5957 | -104.5381
8.5995 | -104.519
8.6032 | -104.4939
8.607 | -104.4685
8.6108 | -104.4386
8.6145 | -104.4083
8.6183 | -104.3771
8.622 | -104.3543
8.6258 | -104.336
8.6296 | -104.3209
8.6333 | -104.3146
8.6371 | -104.3074
8.6408 | -104.3007
8.6446 | -104.294
8.6483 | -104.293
8.6521 | -104.288
8.6559 | -104.2753
8.6596 | -104.2625
8.6634 | -104.2421
8.6671 | -104.2203
8.6709 | -104.1998
8.6747 | -104.1737
8.6784 | -104.1447
8.6822 | -104.1159
8.6859 | -104.0919
8.6897 | -104.0712
8.6934 | -104.0608
8.6972 | -104.0497
8.701 | -104.0378
8.7047 | -104.0307
8.7085 | -104.0273
8.7122 | -104.0288
8.716 | -104.0359
8.7198 | -104.0483
8.7235 | -104.0655
8.7273 | -104.0851
8.731 | -104.1056
8.7348 | -104.1261
8.7385 | -104.1533
8.7423 | -104.1752
8.7461 | -104.1965
8.7498 | -104.2168
8.7536 | -104.2379
8.7573 | -104.2531
8.7611 | -104.2589
8.7649 | -104.2646
8.7686 | -104.2705
8.7724 | -104.2724
8.7761 | -104.2689
8.7799 | -104.2688
8.7836 | -104.2759
8.7874 | -104.2793
8.7912 | -104.2847
8.7949 | -104.2925
8.7987 | -104.298
8.8024 | -104.3073
8.8062 | -104.3228
8.81 | -104.3383
8.8137 | -104.3496
8.8175 | -104.3532
8.8212 | -104.3575
8.825 | -104.3606
8.8288 | -104.3605
8.8325 | -104.3556
8.8363 | -104.3483
8.84 | -104.339
8.8438 | -104.3255
8.8475 | -104.3079
8.8513 | -104.2887
8.8551 | -104.268
8.8588 | -104.2442
8.8626 | -104.2221
8.8663 | -104.2003
8.8701 | -104.175
8.8739 | -104.1557
8.8776 | -104.1399
8.8814 | -104.1269
8.8851 | -104.1181
8.8889 | -104.1163
8.8926 | -104.1232
8.8964 | -104.1343
8.9002 | -104.1453
8.9039 | -104.1646
8.9077 | -104.188
8.9114 | -104.2145
8.9152 | -104.2399
8.919 | -104.2729
8.9227 | -104.3024
8.9265 | -104.325
8.9302 | -104.3498
8.934 | -104.3708
8.9377 | -104.3921
8.9415 | -104.4118
8.9453 | -104.4269
8.949 | -104.4365
8.9528 | -104.4403
8.9565 | -104.4422
8.9603 | -104.4472
8.9641 | -104.4563
8.9678 | -104.4646
8.9716 | -104.4753
8.9753 | -104.4864
8.9791 | -104.5002
8.9829 | -104.518
8.9866 | -104.5363
8.9904 | -104.556
8.9941 | -104.5818
8.9979 | -104.6133
9.0016 | -104.6457
9.0054 | -104.6739
9.0092 | -104.7024
9.0129 | -104.7331
9.0167 | -104.7629
9.0204 | -104.7941
9.0242 | -104.8193
9.028 | -104.8409
9.0317 | -104.8623
9.0355 | -104.8727
9.0392 | -104.8733
9.043 | -104.8653
9.0467 | -104.8571
9.0505 | -104.8387
9.0543 | -104.8179
9.058 | -104.7959
9.0618 | -104.7638
9.0655 | -104.7294
9.0693 | -104.6897
9.0731 | -104.651
9.0768 | -104.6166
9.0806 | -104.5917
9.0843 | -104.5698
9.0881 | -104.5485
9.0918 | -104.5333
9.0956 | -104.5228
9.0994 | -104.5189
9.1031 | -104.5211
9.1069 | -104.5308
9.1106 | -104.5484
9.1144 | -104.5675
9.1182 | -104.5923
9.1219 | -104.6213
9.1257 | -104.649
9.1294 | -104.6786
9.1332 | -104.7133
9.137 | -104.7527
9.1407 | -104.7895
9.1445 | -104.8234
9.1482 | -104.8582
9.152 | -104.8854
9.1557 | -104.9085
9.1595 | -104.9271
9.1633 | -104.9358
9.167 | -104.9404
9.1708 | -104.9408
9.1745 | -104.9362
9.1783 | -104.9193
9.1821 | -104.9014
9.1858 | -104.8791
9.1896 | -104.8529
9.1933 | -104.8294
9.1971 | -104.8087
9.2008 | -104.788
9.2046 | -104.7665
9.2084 | -104.7489
9.2121 | -104.7335
9.2159 | -104.7406
9.2196 | -104.7378
9.2234 | -104.7371
9.2272 | -104.7401
9.2309 | -104.7471
9.2347 | -104.7597
9.2384 | -104.771
9.2422 | -104.7847
9.2459 | -104.7999
9.2497 | -104.8195
9.2535 | -104.8412
9.2572 | -104.8635
9.261 | -104.8857
9.2647 | -104.9071
9.2685 | -104.9284
9.2723 | -104.9506
9.276 | -104.9728
9.2798 | -104.9876
9.2835 | -105.0017
9.2873 | -105.0172
9.2911 | -105.0295
9.2948 | -105.0351
9.2986 | -105.0376
9.3023 | -105.0388
9.3061 | -105.0354
9.3098 | -105.0262
9.3136 | -105.015
9.3174 | -104.9971
9.3211 | -104.9748
9.3249 | -104.954
9.3286 | -104.9285
9.3324 | -104.8977
9.3362 | -104.8732
9.3399 | -104.8496
9.3437 | -104.8225
9.3474 | -104.7919
9.3512 | -104.7647
9.3549 | -104.7393
9.3587 | -104.721
9.3625 | -104.7103
9.3662 | -104.7019
9.37 | -104.6977
9.3737 | -104.696
9.3775 | -104.7034
9.3813 | -104.7133
9.385 | -104.722
9.3888 | -104.7364
9.3925 | -104.7522
9.3963 | -104.7739
9.4 | -104.799
9.4038 | -104.8235
9.4076 | -104.8491
9.4113 | -104.8726
9.4151 | -104.8995
9.4188 | -104.9274
9.4226 | -104.9499
9.4264 | -104.9655
9.4301 | -104.9795
9.4339 | -104.989
9.4376 | -104.9966
9.4414 | -104.9974
9.4451 | -104.996
9.4489 | -104.989
9.4527 | -104.9831
9.4564 | -104.9802
9.4602 | -104.9814
9.4639 | -104.9817
9.4677 | -104.9855
9.4715 | -104.9942
9.4752 | -105.001
9.479 | -105.0122
9.4827 | -105.0277
9.4865 | -105.0508
9.4903 | -105.0726
9.494 | -105.0866
9.4978 | -105.0996
9.5015 | -105.1113
9.5053 | -105.1235
9.509 | -105.1269
9.5128 | -105.1295
9.5166 | -105.1234
9.5203 | -105.1133
9.5241 | -105.0965
9.5278 | -105.0757
9.5316 | -105.051
9.5354 | -105.0177
9.5391 | -104.9824
9.5429 | -104.9486
9.5466 | -104.9114
9.5504 | -104.8698
9.5541 | -104.8292
9.5579 | -104.7958
9.5617 | -104.7661
9.5654 | -104.7438
9.5692 | -104.7237
9.5729 | -104.7106
9.5767 | -104.7002
9.5805 | -104.697
9.5842 | -104.7006
9.588 | -104.7117
9.5917 | -104.7251
9.5955 | -104.7392
9.5992 | -104.7511
9.603 | -104.7616
9.6068 | -104.7794
9.6105 | -104.7972
9.6143 | -104.8116
9.618 | -104.8238
9.6218 | -104.8324
9.6256 | -104.838
9.6293 | -104.8401
9.6331 | -104.8341
9.6368 | -104.8219
9.6406 | -104.8061
9.6443 | -104.7808
9.6481 | -104.7573
9.6519 | -104.7299
9.6556 | -104.7012
9.6594 | -104.6705
9.6631 | -104.644
9.6669 | -104.6113
9.6707 | -104.5868
9.6744 | -104.5648
9.6782 | -104.5426
9.6819 | -104.5317
9.6857 | -104.5272
9.6895 | -104.5315
9.6932 | -104.5348
9.697 | -104.5427
9.7007 | -104.5572
9.7045 | -104.5715
9.7082 | -104.5918
9.712 | -104.6179
9.7158 | -104.6494
9.7195 | -104.6764
9.7233 | -104.6997
9.727 | -104.7204
9.7308 | -104.7348
9.7346 | -104.7476
9.7383 | -104.7606
9.7421 | -104.7704
9.7458 | -104.7701
9.7496 | -104.7665
9.7533 | -104.7604
9.7571 | -104.7504
9.7609 | -104.7447
9.7646 | -104.7384
9.7684 | -104.7337
9.7721 | -104.7318
9.7759 | -104.7379
9.7797 | -104.7431
9.7834 | -104.7465
9.7872 | -104.7521
9.7909 | -104.7627
9.7947 | -104.7759
9.7985 | -104.7955
9.8022 | -104.8195
9.806 | -104.8442
9.8097 | -104.8771
9.8135 | -104.9102
9.8172 | -104.9517
9.821 | -104.9981
9.8248 | -105.037
9.8285 | -105.0737
9.8323 | -105.113
9.836 | -105.1507
9.8398 | -105.1812
9.8436 | -105.21
9.8473 | -105.237
9.8511 | -105.2544
9.8548 | -105.2599
9.8586 | -105.261
9.8623 | -105.2619
9.8661 | -105.2567
9.8699 | -105.2464
9.8736 | -105.2338
9.8774 | -105.2142
9.8811 | -105.1893
9.8849 | -105.1607
9.8887 | -105.1333
9.8924 | -105.1114
9.8962 | -105.0913
9.8999 | -105.0707
9.9037 | -105.0529
9.9074 | -105.0379
9.9112 | -105.0254
9.915 | -105.0167
9.9187 | -105.0124
9.9225 | -105.0072
9.9262 | -105.0013
9.93 | -105.0012
9.9338 | -105.0065
9.9375 | -105.0087
9.9413 | -105.0093
9.945 | -105.0069
9.9488 | -105.0058
9.9525 | -105.0087
9.9563 | -105.0067
9.9601 | -105.0024
9.9638 | -104.9991
9.9676 | -104.9934
9.9713 | -104.9903
9.9751 | -104.9916
9.9789 | -104.9864
9.9826 | -104.9817
9.9864 | -104.9767
9.9901 | -104.9665
9.9939 | -104.9552
9.9977 | -104.9449
10.0014 | -104.9345
10.0052 | -104.9233
10.0089 | -104.9135
10.0127 | -104.9042
10.0164 | -104.8976
10.0202 | -104.8947
10.024 | -104.8931
10.0277 | -104.894
10.0315 | -104.8993
10.0352 | -104.9004
10.039 | -104.8986
10.0428 | -104.9027
10.0465 | -104.9069
10.0503 | -104.9138
10.054 | -104.9226
10.0578 | -104.9281
10.0615 | -104.9284
10.0653 | -104.9286
10.0691 | -104.9295
10.0728 | -104.9301
10.0766 | -104.928
10.0803 | -104.9207
10.0841 | -104.917
10.0879 | -104.9127
10.0916 | -104.9098
10.0954 | -104.9077
10.0991 | -104.9026
10.1029 | -104.8998
10.1066 | -104.901
10.1104 | -104.9021
10.1142 | -104.907
10.1179 | -104.9169
10.1217 | -104.9289
10.1254 | -104.9422
10.1292 | -104.9596
10.133 | -104.9827
10.1367 | -105.0118
10.1405 | -105.0417
10.1442 | -105.0764
10.148 | -105.1172
10.1517 | -105.1565
10.1555 | -105.1949
10.1593 | -105.2342
10.163 | -105.2769
10.1668 | -105.3176
10.1705 | -105.3571
10.1743 | -105.3881
10.1781 | -105.4188
10.1818 | -105.4499
10.1856 | -105.477
10.1893 | -105.4966
10.1931 | -105.5111
10.1969 | -105.5147
10.2006 | -105.5145
10.2044 | -105.5143
10.2081 | -105.5094
10.2119 | -105.5017
10.2156 | -105.4915
10.2194 | -105.4761
10.2232 | -105.4593
10.2269 | -105.4442
10.2307 | -105.4235
10.2344 | -105.4071
10.2382 | -105.3936
10.242 | -105.379
10.2457 | -105.3718
10.2495 | -105.3681
10.2532 | -105.3639
10.257 | -105.3661
10.2607 | -105.3756
10.2645 | -105.3888
10.2683 | -105.4078
10.272 | -105.4353
10.2758 | -105.4641
10.2795 | -105.4937
10.2833 | -105.5229
10.2871 | -105.5607
10.2908 | -105.5961
10.2946 | -105.6277
10.2983 | -105.662
10.3021 | -105.6903
10.3058 | -105.7237
10.3096 | -105.7574
10.3134 | -105.7788
10.3171 | -105.7991
10.3209 | -105.8153
10.3246 | -105.8312
10.3284 | -105.8395
10.3322 | -105.8433
10.3359 | -105.8409
10.3397 | -105.8297
10.3434 | -105.8179
10.3472 | -105.8
10.351 | -105.773
10.3547 | -105.7424
10.3585 | -105.7108
10.3622 | -105.6693
10.366 | -105.6237
10.3697 | -105.5817
10.3735 | -105.5388
10.3773 | -105.4969
10.381 | -105.457
10.3848 | -105.4196
10.3885 | -105.3815
10.3923 | -105.3456
10.3961 | -105.3145
10.3998 | -105.2924
10.4036 | -105.2789
10.4073 | -105.2663
10.4111 | -105.2565
10.4148 | -105.2513
10.4186 | -105.2495
10.4224 | -105.2537
10.4261 | -105.2649
10.4299 | -105.2836
10.4336 | -105.3083
10.4374 | -105.3337
10.4412 | -105.3619
10.4449 | -105.3957
10.4487 | -105.4289
10.4524 | -105.4591
10.4562 | -105.4882
10.4599 | -105.5174
10.4637 | -105.546
10.4675 | -105.5756
10.4712 | -105.5944
10.475 | -105.6114
10.4787 | -105.6227
10.4825 | -105.6287
10.4863 | -105.6294
10.49 | -105.6269
10.4938 | -105.6239
10.4975 | -105.61
10.5013 | -105.5973
10.5051 | -105.5813
10.5088 | -105.5632
10.5126 | -105.5415
10.5163 | -105.5243
10.5201 | -105.5065
10.5238 | -105.4824
10.5276 | -105.4611
10.5314 | -105.4396
10.5351 | -105.4227
10.5389 | -105.4062
10.5426 | -105.3972
10.5464 | -105.3921
10.5502 | -105.386
10.5539 | -105.3796
10.5577 | -105.3751
10.5614 | -105.3797
10.5652 | -105.3894
10.5689 | -105.3998
10.5727 | -105.4157
10.5765 | -105.4379
10.5802 | -105.4574
10.584 | -105.4772
10.5877 | -105.4989
10.5915 | -105.517
10.5953 | -105.5357
10.599 | -105.555
10.6028 | -105.571
10.6065 | -105.5857
10.6103 | -105.5954
10.614 | -105.5969
10.6178 | -105.5974
10.6216 | -105.5934
10.6253 | -105.586
10.6291 | -105.5721
10.6328 | -105.5527
10.6366 | -105.5283
10.6404 | -105.5015
10.6441 | -105.4734
10.6479 | -105.442
10.6516 | -105.4112
10.6554 | -105.3783
10.6592 | -105.3456
10.6629 | -105.3121
10.6667 | -105.2816
10.6704 | -105.2506
10.6742 | -105.2224
10.6779 | -105.1994
10.6817 | -105.1777
10.6855 | -105.1617
10.6892 | -105.1471
10.693 | -105.1368
10.6967 | -105.1259
10.7005 | -105.1194
10.7043 | -105.1159
10.708 | -105.1182
10.7118 | -105.1225
10.7155 | -105.1238
10.7193 | -105.1249
10.723 | -105.1327
10.7268 | -105.141
10.7306 | -105.148
10.7343 | -105.1575
10.7381 | -105.1648
10.7418 | -105.1724
10.7456 | -105.1872
10.7494 | -105.1988
10.7531 | -105.2038
10.7569 | -105.2082
10.7606 | -105.2127
10.7644 | -105.2192
10.7681 | -105.2238
10.7719 | -105.221
10.7757 | -105.2223
10.7794 | -105.2269
10.7832 | -105.2263
10.7869 | -105.2255
10.7907 | -105.2253
10.7945 | -105.2248
10.7982 | -105.2223
10.802 | -105.2171
10.8057 | -105.2117
10.8095 | -105.206
10.8132 | -105.1979
10.817 | -105.1903
10.8208 | -105.1821
10.8245 | -105.1768
10.8283 | -105.1702
10.832 | -105.1649
10.8358 | -105.1597
10.8396 | -105.153
10.8433 | -105.1482
10.8471 | -105.1431
10.8508 | -105.1413
10.8546 | -105.1424
10.8584 | -105.1414
10.8621 | -105.1413
10.8659 | -105.1396
10.8696 | -105.1388
10.8734 | -105.1399
10.8771 | -105.141
10.8809 | -105.1366
10.8847 | -105.1315
10.8884 | -105.1241
10.8922 | -105.1173
10.8959 | -105.1125
10.8997 | -105.1056
10.9035 | -105.0986
10.9072 | -105.0912
10.911 | -105.0849
10.9147 | -105.082
10.9185 | -105.0746
10.9222 | -105.0657
10.926 | -105.0582
10.9298 | -105.0523
10.9335 | -105.0448
10.9373 | -105.038
10.941 | -105.0337
10.9448 | -105.0273
10.9486 | -105.0215
10.9523 | -105.0161
10.9561 | -105.0146
10.9598 | -105.0181
10.9636 | -105.0207
10.9673 | -105.0225
10.9711 | -105.026
10.9749 | -105.0284
10.9786 | -105.0302
10.9824 | -105.0304
10.9861 | -105.0374
10.9899 | -105.044
10.9937 | -105.0468
10.9974 | -105.0475
11.0012 | -105.0529
11.0049 | -105.0643
11.0087 | -105.0733
11.0124 | -105.0802
11.0162 | -105.0882
11.02 | -105.0972
11.0237 | -105.106
11.0275 | -105.1164
11.0312 | -105.1267
11.035 | -105.1334
11.0388 | -105.137
11.0425 | -105.1437
11.0463 | -105.1515
11.05 | -105.157
11.0538 | -105.1605
11.0576 | -105.1606
11.0613 | -105.1604
11.0651 | -105.1605
11.0688 | -105.1623
11.0726 | -105.163
11.0763 | -105.1631
11.0801 | -105.1658
11.0839 | -105.1676
11.0876 | -105.1711
11.0914 | -105.1743
11.0951 | -105.178
11.0989 | -105.183
11.1027 | -105.1876
11.1064 | -105.1947
11.1102 | -105.1992
11.1139 | -105.2018
11.1177 | -105.202
11.1214 | -105.204
11.1252 | -105.2108
11.129 | -105.2171
11.1327 | -105.2214
11.1365 | -105.2266
11.1402 | -105.2321
11.144 | -105.2332
11.1478 | -105.2343
11.1515 | -105.2398
11.1553 | -105.2487
11.159 | -105.2591
11.1628 | -105.2665
11.1665 | -105.2694
11.1703 | -105.2723
11.1741 | -105.2781
11.1778 | -105.2796
11.1816 | -105.2786
11.1853 | -105.2854
11.1891 | -105.2885
11.1929 | -105.2909
11.1966 | -105.2941
11.2004 | -105.298
11.2041 | -105.2943
11.2079 | -105.2871
11.2117 | -105.2864
11.2154 | -105.289
11.2192 | -105.2896
11.2229 | -105.2872
11.2267 | -105.2826
11.2304 | -105.2824
11.2342 | -105.2861
11.238 | -105.2893
11.2417 | -105.293
11.2455 | -105.2989
11.2492 | -105.3058
11.253 | -105.3158
11.2568 | -105.3258
11.2605 | -105.3349
11.2643 | -105.347
11.268 | -105.3582
11.2718 | -105.3717
11.2755 | -105.3889
11.2793 | -105.4029
11.2831 | -105.4112
11.2868 | -105.4191
11.2906 | -105.4309
11.2943 | -105.4398
11.2981 | -105.4455
11.3019 | -105.4477
11.3056 | -105.4512
11.3094 | -105.4473
11.3131 | -105.439
11.3169 | -105.4282
11.3206 | -105.4178
11.3244 | -105.402
11.3282 | -105.3863
11.3319 | -105.3707
11.3357 | -105.3517
11.3394 | -105.3356
11.3432 | -105.3233
11.347 | -105.3048
11.3507 | -105.2892
11.3545 | -105.2774
11.3582 | -105.2712
11.362 | -105.2684
11.3658 | -105.2695
11.3695 | -105.2739
11.3733 | -105.2785
11.377 | -105.2838
11.3808 | -105.2945
11.3845 | -105.3151
11.3883 | -105.3338
11.3921 | -105.3508
11.3958 | -105.3657
11.3996 | -105.3834
11.4033 | -105.4015
11.4071 | -105.4197
11.4109 | -105.4418
11.4146 | -105.4648
11.4184 | -105.4853
11.4221 | -105.5057
11.4259 | -105.5271
11.4296 | -105.5467
11.4334 | -105.5613
11.4372 | -105.5696
11.4409 | -105.5809
11.4447 | -105.5929
11.4484 | -105.6041
11.4522 | -105.6082
11.456 | -105.6065
11.4597 | -105.6037
11.4635 | -105.5987
11.4672 | -105.5921
11.471 | -105.581
11.4747 | -105.5659
11.4785 | -105.5471
11.4823 | -105.5297
11.486 | -105.5076
11.4898 | -105.487
11.4935 | -105.4657
11.4973 | -105.4441
11.5011 | -105.4223
11.5048 | -105.4008
11.5086 | -105.3811
11.5123 | -105.3612
11.5161 | -105.3467
11.5198 | -105.3368
11.5236 | -105.3284
11.5274 | -105.3201
11.5311 | -105.316
11.5349 | -105.3157
11.5386 | -105.3189
11.5424 | -105.3284
11.5462 | -105.3416
11.5499 | -105.3577
11.5537 | -105.3731
11.5574 | -105.391
11.5612 | -105.4108
11.565 | -105.433
11.5687 | -105.4511
11.5725 | -105.4713
11.5762 | -105.4925
11.58 | -105.5068
11.5837 | -105.5203
11.5875 | -105.5338
11.5913 | -105.5452
11.595 | -105.5546
11.5988 | -105.5629
11.6025 | -105.5722
11.6063 | -105.5813
11.6101 | -105.5809
11.6138 | -105.5766
11.6176 | -105.5675
11.6213 | -105.5584
11.6251 | -105.545
11.6288 | -105.5367
11.6326 | -105.5254
11.6364 | -105.5135
11.6401 | -105.5031
11.6439 | -105.4935
11.6476 | -105.4847
11.6514 | -105.473
11.6552 | -105.4649
11.6589 | -105.4587
11.6627 | -105.4545
11.6664 | -105.4547
11.6702 | -105.4543
11.6739 | -105.4563
11.6777 | -105.4575
11.6815 | -105.4612
11.6852 | -105.4662
11.689 | -105.4723
11.6927 | -105.4794
11.6965 | -105.4906
11.7003 | -105.5018
11.704 | -105.5127
11.7078 | -105.5186
11.7115 | -105.524
11.7153 | -105.5298
11.7191 | -105.5381
11.7228 | -105.548
11.7266 | -105.5575
11.7303 | -105.5685
11.7341 | -105.5802
11.7378 | -105.5921
11.7416 | -105.601
11.7454 | -105.6103
11.7491 | -105.6183
11.7529 | -105.6211
11.7566 | -105.6254
11.7604 | -105.6321
11.7642 | -105.6366
11.7679 | -105.6401
11.7717 | -105.6406
11.7754 | -105.6377
11.7792 | -105.6376
11.7829 | -105.6386
11.7867 | -105.6394
11.7905 | -105.6377
11.7942 | -105.6347
11.798 | -105.6238
11.8017 | -105.6145
11.8055 | -105.6045
11.8093 | -105.5905
11.813 | -105.575
11.8168 | -105.5619
11.8205 | -105.5503
11.8243 | -105.5357
11.828 | -105.5171
11.8318 | -105.5014
11.8356 | -105.4878
11.8393 | -105.4809
11.8431 | -105.4763
11.8468 | -105.4737
11.8506 | -105.472
11.8544 | -105.4716
11.8581 | -105.4735
11.8619 | -105.4783
11.8656 | -105.4826
11.8694 | -105.4898
11.8732 | -105.5032
11.8769 | -105.524
11.8807 | -105.5456
11.8844 | -105.5701
11.8882 | -105.5921
11.8919 | -105.6104
11.8957 | -105.6313
11.8995 | -105.6569
11.9032 | -105.6836
11.907 | -105.7095
11.9107 | -105.735
11.9145 | -105.7582
11.9183 | -105.7827
11.922 | -105.8086
11.9258 | -105.8195
11.9295 | -105.8369
11.9333 | -105.8492
11.937 | -105.8611
11.9408 | -105.8726
11.9446 | -105.8847
11.9483 | -105.894
11.9521 | -105.8946
11.9558 | -105.8983
11.9596 | -105.9034
11.9634 | -105.907
11.9671 | -105.9108
11.9709 | -105.913
11.9746 | -105.9116
11.9784 | -105.9106
11.9821 | -105.909
11.9859 | -105.9035
11.9897 | -105.8963
11.9934 | -105.8907
11.9972 | -105.8871
12.0009 | -105.8851
12.0047 | -105.8792
12.0085 | -105.869
12.0122 | -105.8579
12.016 | -105.8445
12.0197 | -105.8299
12.0235 | -105.8174
12.0272 | -105.8009
12.031 | -105.7823
12.0348 | -105.7641
12.0385 | -105.7489
12.0423 | -105.7319
12.046 | -105.7103
12.0498 | -105.6884
12.0536 | -105.6669
12.0573 | -105.6478
12.0611 | -105.6314
12.0648 | -105.6168
12.0686 | -105.6027
12.0724 | -105.5853
12.0761 | -105.5692
12.0799 | -105.5592
12.0836 | -105.551
12.0874 | -105.5413
12.0911 | -105.5317
12.0949 | -105.5213
12.0987 | -105.5143
12.1024 | -105.5094
12.1062 | -105.5068
12.1099 | -105.5066
12.1137 | -105.5028
12.1175 | -105.4988
12.1212 | -105.4971
12.125 | -105.5002
12.1287 | -105.5065
12.1325 | -105.5081
12.1362 | -105.5116
12.14 | -105.5186
12.1438 | -105.5275
12.1475 | -105.5374
12.1513 | -105.551
12.155 | -105.5659
12.1588 | -105.5799
12.1626 | -105.5971
12.1663 | -105.6115
12.1701 | -105.6236
12.1738 | -105.6355
12.1776 | -105.6425
12.1813 | -105.6447
12.1851 | -105.6499
12.1889 | -105.651
12.1926 | -105.6503
12.1964 | -105.6477
12.2001 | -105.6444
12.2039 | -105.6402
12.2077 | -105.6324
12.2114 | -105.6223
12.2152 | -105.6081
12.2189 | -105.5921
12.2227 | -105.5762
12.2265 | -105.5599
12.2302 | -105.5438
12.234 | -105.5267
12.2377 | -105.5106
12.2415 | -105.4976
12.2452 | -105.4842
12.249 | -105.4759
12.2528 | -105.4675
12.2565 | -105.4592
12.2603 | -105.4553
12.264 | -105.4556
12.2678 | -105.4585
12.2716 | -105.4652
12.2753 | -105.4738
12.2791 | -105.4831
12.2828 | -105.4974
12.2866 | -105.5091
12.2903 | -105.5249
12.2941 | -105.5422
12.2979 | -105.56
12.3016 | -105.577
12.3054 | -105.5963
12.3091 | -105.6144
12.3129 | -105.6272
12.3167 | -105.6355
12.3204 | -105.6421
12.3242 | -105.6484
12.3279 | -105.6529
12.3317 | -105.653
12.3354 | -105.6539
12.3392 | -105.649
12.343 | -105.6409
12.3467 | -105.6315
12.3505 | -105.619
12.3542 | -105.609
12.358 | -105.5996
12.3618 | -105.5856
12.3655 | -105.574
12.3693 | -105.5637
12.373 | -105.5538
12.3768 | -105.5421
12.3805 | -105.5346
12.3843 | -105.5292
12.3881 | -105.5269
12.3918 | -105.5287
12.3956 | -105.5323
12.3993 | -105.5386
12.4031 | -105.5485
12.4069 | -105.5565
12.4106 | -105.5646
12.4144 | -105.5735
12.4181 | -105.5818
12.4219 | -105.5909
12.4257 | -105.5992
12.4294 | -105.608
12.4332 | -105.6171
12.4369 | -105.6239
12.4407 | -105.6301
12.4444 | -105.6337
12.4482 | -105.6337
12.452 | -105.6304
12.4557 | -105.6266
12.4595 | -105.6208
12.4632 | -105.6192
12.467 | -105.6177
12.4708 | -105.6154
12.4745 | -105.6157
12.4783 | -105.6171
12.482 | -105.62
12.4858 | -105.6235
12.4895 | -105.6333
12.4933 | -105.6472
12.4971 | -105.6619
12.5008 | -105.6784
12.5046 | -105.6955
12.5083 | -105.7129
12.5121 | -105.7337
12.5159 | -105.7549
12.5196 | -105.7793
12.5234 | -105.8048
12.5271 | -105.8315
12.5309 | -105.8604
12.5346 | -105.8894
12.5384 | -105.9171
12.5422 | -105.9465
12.5459 | -105.974
12.5497 | -105.9977
12.5534 | -106.0214
12.5572 | -106.0432
12.561 | -106.0629
12.5647 | -106.0815
12.5685 | -106.0964
12.5722 | -106.1075
12.576 | -106.1152
12.5798 | -106.12
12.5835 | -106.1242
12.5873 | -106.1314
12.591 | -106.135
12.5948 | -106.1354
12.5985 | -106.1366
12.6023 | -106.1377
12.6061 | -106.1344
12.6098 | -106.1315
12.6136 | -106.1291
12.6173 | -106.1265
12.6211 | -106.1236
12.6249 | -106.1218
12.6286 | -106.121
12.6324 | -106.121
12.6361 | -106.1193
12.6399 | -106.1189
12.6436 | -106.1166
12.6474 | -106.1137
12.6512 | -106.1081
12.6549 | -106.1051
12.6587 | -106.1012
12.6624 | -106.0966
12.6662 | -106.0901
12.67 | -106.0821
12.6737 | -106.0729
12.6775 | -106.0669
12.6812 | -106.0632
12.685 | -106.058
12.6887 | -106.0538
12.6925 | -106.049
12.6963 | -106.0474
12.7 | -106.049
12.7038 | -106.0503
12.7075 | -106.0526
12.7113 | -106.0543
12.7151 | -106.0563
12.7188 | -106.0602
12.7226 | -106.0671
12.7263 | -106.0714
12.7301 | -106.0781
12.7339 | -106.086
12.7376 | -106.0925
12.7414 | -106.1013
12.7451 | -106.1097
12.7489 | -106.1186
12.7526 | -106.128
12.7564 | -106.1372
12.7602 | -106.146
12.7639 | -106.1517
12.7677 | -106.1534
12.7714 | -106.1509
12.7752 | -106.1531
12.779 | -106.1569
12.7827 | -106.1561
12.7865 | -106.1531
12.7902 | -106.1507
12.794 | -106.1474
12.7977 | -106.1422
12.8015 | -106.1381
12.8053 | -106.1353
12.809 | -106.1297
12.8128 | -106.1244
12.8165 | -106.1171
12.8203 | -106.1081
12.8241 | -106.0984
12.8278 | -106.0905
12.8316 | -106.083
12.8353 | -106.0782
12.8391 | -106.0712
12.8428 | -106.0647
12.8466 | -106.0595
12.8504 | -106.0556
12.8541 | -106.0498
12.8579 | -106.0466
12.8616 | -106.041
12.8654 | -106.0376
12.8692 | -106.0362
12.8729 | -106.0354
12.8767 | -106.0291
12.8804 | -106.0213
12.8842 | -106.0155
12.8879 | -106.0099
12.8917 | -106.0038
12.8955 | -105.9943
12.8992 | -105.9851
12.903 | -105.9767
12.9067 | -105.9689
12.9105 | -105.962
12.9143 | -105.9513
12.918 | -105.9374
12.9218 | -105.9278
12.9255 | -105.9165
12.9293 | -105.9035
12.9331 | -105.8914
12.9368 | -105.8773
12.9406 | -105.8649
12.9443 | -105.8552
12.9481 | -105.8466
12.9518 | -105.8401
12.9556 | -105.8359
12.9594 | -105.8295
12.9631 | -105.823
12.9669 | -105.8192
12.9706 | -105.8181
12.9744 | -105.8183
12.9782 | -105.8168
12.9819 | -105.8159
12.9857 | -105.8123
12.9894 | -105.8103
12.9932 | -105.8131
12.9969 | -105.8131
13.0007 | -105.8118
13.0045 | -105.8093
13.0082 | -105.8068
13.012 | -105.8062
13.0157 | -105.8054
13.0195 | -105.806
13.0233 | -105.8069
13.027 | -105.8074
13.0308 | -105.8074
13.0345 | -105.8088
13.0383 | -105.8144
13.042 | -105.8201
13.0458 | -105.8217
13.0496 | -105.8256
13.0533 | -105.8291
13.0571 | -105.8351
13.0608 | -105.8413
13.0646 | -105.8519
13.0684 | -105.8625
13.0721 | -105.8721
13.0759 | -105.8837
13.0796 | -105.8977
13.0834 | -105.9131
13.0872 | -105.9276
13.0909 | -105.9403
13.0947 | -105.9526
13.0984 | -105.964
13.1022 | -105.9741
13.1059 | -105.99
13.1097 | -106.0024
13.1135 | -106.012
13.1172 | -106.0233
13.121 | -106.036
13.1247 | -106.0505
13.1285 | -106.06
13.1323 | -106.0667
13.136 | -106.0725
13.1398 | -106.0734
13.1435 | -106.079
13.1473 | -106.0824
13.151 | -106.0838
13.1548 | -106.0835
13.1586 | -106.0824
13.1623 | -106.079
13.1661 | -106.0806
13.1698 | -106.0817
13.1736 | -106.0814
13.1774 | -106.0792
13.1811 | -106.0757
13.1849 | -106.0734
13.1886 | -106.075
13.1924 | -106.0764
13.1961 | -106.0763
13.1999 | -106.0729
13.2037 | -106.0713
13.2074 | -106.0713
13.2112 | -106.0707
13.2149 | -106.0659
13.2187 | -106.0609
13.2225 | -106.0544
13.2262 | -106.0464
13.23 | -106.0394
13.2337 | -106.0343
13.2375 | -106.0294
13.2413 | -106.0213
13.245 | -106.011
13.2488 | -105.9983
13.2525 | -105.9865
13.2563 | -105.9737
13.26 | -105.962
13.2638 | -105.951
13.2676 | -105.9415
13.2713 | -105.9339
13.2751 | -105.9327
13.2788 | -105.9333
13.2826 | -105.9326
13.2864 | -105.9326
13.2901 | -105.9355
13.2939 | -105.9444
13.2976 | -105.9518
13.3014 | -105.9602
13.3051 | -105.9712
13.3089 | -105.9808
13.3127 | -105.9898
13.3164 | -105.999
13.3202 | -106.0141
13.3239 | -106.0328
13.3277 | -106.0521
13.3315 | -106.0712
13.3352 | -106.0815
13.339 | -106.0947
13.3427 | -106.1092
13.3465 | -106.1227
13.3502 | -106.1361
13.354 | -106.1507
13.3578 | -106.1574
13.3615 | -106.1643
13.3653 | -106.1717
13.369 | -106.1756
13.3728 | -106.1762
13.3766 | -106.1763
13.3803 | -106.1738
13.3841 | -106.1704
13.3878 | -106.1667
13.3916 | -106.1635
13.3953 | -106.1556
13.3991 | -106.1468
13.4029 | -106.1351
13.4066 | -106.1221
13.4104 | -106.1051
13.4141 | -106.0873
13.4179 | -106.0694
13.4217 | -106.05
13.4254 | -106.0308
13.4292 | -106.0095
13.4329 | -105.9855
13.4367 | -105.9607
13.4405 | -105.9383
13.4442 | -105.9143
13.448 | -105.8932
13.4517 | -105.8658
13.4555 | -105.8326
13.4592 | -105.8012
13.463 | -105.7734
13.4668 | -105.7446
13.4705 | -105.7163
13.4743 | -105.6939
13.478 | -105.6716
13.4818 | -105.6535
13.4856 | -105.6363
13.4893 | -105.6189
13.4931 | -105.6038
13.4968 | -105.5938
13.5006 | -105.5898
13.5043 | -105.5897
13.5081 | -105.5936
13.5119 | -105.5993
13.5156 | -105.6042
13.5194 | -105.6136
13.5231 | -105.6301
13.5269 | -105.6536
13.5307 | -105.6775
13.5344 | -105.7022
13.5382 | -105.7287
13.5419 | -105.7577
13.5457 | -105.7916
13.5494 | -105.828
13.5532 | -105.8632
13.557 | -105.8961
13.5607 | -105.9299
13.5645 | -105.965
13.5682 | -106.0001
13.572 | -106.0319
13.5758 | -106.0614
13.5795 | -106.0912
13.5833 | -106.1187
13.587 | -106.1449
13.5908 | -106.1741
13.5945 | -106.1974
13.5983 | -106.2165
13.6021 | -106.2375
13.6058 | -106.2586
13.6096 | -106.2751
13.6133 | -106.2889
13.6171 | -106.297
13.6209 | -106.303
13.6246 | -106.3095
13.6284 | -106.3159
13.6321 | -106.3194
13.6359 | -106.3178
13.6397 | -106.3175
13.6434 | -106.3157
13.6472 | -106.3134
13.6509 | -106.3136
13.6547 | -106.3085
13.6584 | -106.2976
13.6622 | -106.2872
13.666 | -106.2765
13.6697 | -106.2653
13.6735 | -106.2576
13.6772 | -106.247
13.681 | -106.2377
13.6848 | -106.2324
13.6885 | -106.2283
13.6923 | -106.2245
13.696 | -106.2196
13.6998 | -106.2141
13.7035 | -106.213
13.7073 | -106.217
13.7111 | -106.2251
13.7148 | -106.2301
13.7186 | -106.2297
13.7223 | -106.2307
13.7261 | -106.236
13.7299 | -106.2433
13.7336 | -106.2496
13.7374 | -106.2548
13.7411 | -106.2591
13.7449 | -106.266
13.7486 | -106.2746
13.7524 | -106.2822
13.7562 | -106.2885
13.7599 | -106.2972
13.7637 | -106.3053
13.7674 | -106.3105
13.7712 | -106.3165
13.775 | -106.3246
13.7787 | -106.3349
13.7825 | -106.3439
13.7862 | -106.3504
13.79 | -106.3551
13.7938 | -106.3644
13.7975 | -106.3752
13.8013 | -106.3837
13.805 | -106.3899
13.8088 | -106.3948
13.8125 | -106.3922
13.8163 | -106.3895
13.8201 | -106.3878
13.8238 | -106.3816
13.8276 | -106.3743
13.8313 | -106.3695
13.8351 | -106.3632
13.8389 | -106.3541
13.8426 | -106.3458
13.8464 | -106.3383
13.8501 | -106.3289
13.8539 | -106.3193
13.8576 | -106.3087
13.8614 | -106.2961
13.8652 | -106.2836
13.8689 | -106.2679
13.8727 | -106.2518
13.8764 | -106.2363
13.8802 | -106.2213
13.884 | -106.2088
13.8877 | -106.1969
13.8915 | -106.1843
13.8952 | -106.1746
13.899 | -106.1675
13.9027 | -106.1603
13.9065 | -106.1545
13.9103 | -106.1502
13.914 | -106.1513
13.9178 | -106.159
13.9215 | -106.1642
13.9253 | -106.1697
13.9291 | -106.1778
13.9328 | -106.1911
13.9366 | -106.2039
13.9403 | -106.219
13.9441 | -106.2365
13.9479 | -106.255
13.9516 | -106.2748
13.9554 | -106.2956
13.9591 | -106.3111
13.9629 | -106.3135
13.9666 | -106.3157
13.9704 | -106.3271
13.9742 | -106.3383
13.9779 | -106.3457
13.9817 | -106.3523
13.9854 | -106.3573
13.9892 | -106.36
13.993 | -106.3567
13.9967 | -106.351
14.0005 | -106.3388
14.0042 | -106.3234
14.008 | -106.304
14.0117 | -106.2817
14.0155 | -106.2557
14.0193 | -106.2267
14.023 | -106.1971
14.0268 | -106.1648
14.0305 | -106.1292
14.0343 | -106.0928
14.0381 | -106.0573
14.0418 | -106.0218
14.0456 | -105.9919
14.0493 | -105.9659
14.0531 | -105.9406
14.0568 | -105.9208
14.0606 | -105.9061
14.0644 | -105.8966
14.0681 | -105.8926
14.0719 | -105.8897
14.0756 | -105.8915
14.0794 | -105.8978
14.0832 | -105.9091
14.0869 | -105.9253
14.0907 | -105.9455
14.0944 | -105.9781
14.0982 | -106.0009
14.102 | -106.0264
14.1057 | -106.0503
14.1095 | -106.0765
14.1132 | -106.1039
14.117 | -106.1304
14.1207 | -106.1569
14.1245 | -106.1864
14.1283 | -106.213
14.132 | -106.2378
14.1358 | -106.2595
14.1395 | -106.281
14.1433 | -106.2978
14.1471 | -106.3128
14.1508 | -106.3268
14.1546 | -106.3422
14.1583 | -106.358
14.1621 | -106.3737
14.1658 | -106.3996
14.1696 | -106.4202
14.1734 | -106.4246
14.1771 | -106.4254
14.1809 | -106.4271
14.1846 | -106.4307
14.1884 | -106.4357
14.1922 | -106.4395
14.1959 | -106.4486
14.1997 | -106.4583
14.2034 | -106.4685
14.2072 | -106.4793
14.2109 | -106.4906
14.2147 | -106.5024
14.2185 | -106.5144
14.2222 | -106.5252
14.226 | -106.5349
14.2297 | -106.545
14.2335 | -106.5538
14.2373 | -106.564
14.241 | -106.5723
14.2448 | -106.5816
14.2485 | -106.5881
14.2523 | -106.5919
14.256 | -106.6002
14.2598 | -106.6046
14.2636 | -106.609
14.2673 | -106.6113
14.2711 | -106.6132
14.2748 | -106.6157
14.2786 | -106.6163
14.2824 | -106.6131
14.2861 | -106.6079
14.2899 | -106.6047
14.2936 | -106.6013
14.2974 | -106.5975
14.3012 | -106.5938
14.3049 | -106.5865
14.3087 | -106.5774
14.3124 | -106.5724
14.3162 | -106.5667
14.3199 | -106.5588
14.3237 | -106.551
14.3275 | -106.5433
14.3312 | -106.5355
14.335 | -106.5283
14.3387 | -106.5224
14.3425 | -106.5194
14.3463 | -106.5161
14.35 | -106.5131
14.3538 | -106.5094
14.3575 | -106.5064
14.3613 | -106.5005
14.365 | -106.4947
14.3688 | -106.4934
14.3726 | -106.4921
14.3763 | -106.4925
14.3801 | -106.4933
14.3838 | -106.493
14.3876 | -106.4952
14.3914 | -106.4971
14.3951 | -106.4998
14.3989 | -106.5021
14.4026 | -106.5031
14.4064 | -106.5062
14.4101 | -106.5097
14.4139 | -106.5124
14.4177 | -106.5145
14.4214 | -106.5174
14.4252 | -106.5201
14.4289 | -106.5231
14.4327 | -106.5287
14.4365 | -106.5324
14.4402 | -106.5373
14.444 | -106.5397
14.4477 | -106.5423
14.4515 | -106.5436
14.4552 | -106.5471
14.459 | -106.5513
14.4628 | -106.5533
14.4665 | -106.5542
14.4703 | -106.5545
14.474 | -106.5552
14.4778 | -106.555
14.4816 | -106.552
14.4853 | -106.5479
14.4891 | -106.5455
14.4928 | -106.5383
14.4966 | -106.5292
14.5004 | -106.5198
14.5041 | -106.5078
14.5079 | -106.4931
14.5116 | -106.4788
14.5154 | -106.4651
14.5191 | -106.4475
14.5229 | -106.4265
14.5267 | -106.4058
14.5304 | -106.3863
14.5342 | -106.3675
14.5379 | -106.3478
14.5417 | -106.3307
14.5455 | -106.3162
14.5492 | -106.2994
14.553 | -106.2848
14.5567 | -106.2735
14.5605 | -106.2618
14.5642 | -106.2506
14.568 | -106.2425
14.5718 | -106.2377
14.5755 | -106.2341
14.5793 | -106.2304
14.583 | -106.2297
14.5868 | -106.232
14.5906 | -106.2328
14.5943 | -106.2319
14.5981 | -106.2355
14.6018 | -106.2371
14.6056 | -106.241
14.6093 | -106.2433
14.6131 | -106.243
14.6169 | -106.243
14.6206 | -106.2449
14.6244 | -106.2466
14.6281 | -106.2476
14.6319 | -106.2449
14.6357 | -106.2382
14.6394 | -106.2301
14.6432 | -106.2249
14.6469 | -106.2218
14.6507 | -106.2158
14.6545 | -106.2124
14.6582 | -106.2093
14.662 | -106.2062
14.6657 | -106.2014
14.6695 | -106.197
14.6732 | -106.1931
14.677 | -106.1924
14.6808 | -106.1903
14.6845 | -106.1898
14.6883 | -106.1908
14.692 | -106.1924
14.6958 | -106.1926
14.6996 | -106.1938
14.7033 | -106.1999
14.7071 | -106.2031
14.7108 | -106.2072
14.7146 | -106.214
14.7183 | -106.2212
14.7221 | -106.2282
14.7259 | -106.2349
14.7296 | -106.2442
14.7334 | -106.2503
14.7371 | -106.2601
14.7409 | -106.2687
14.7447 | -106.2759
14.7484 | -106.2843
14.7522 | -106.2944
14.7559 | -106.3035
14.7597 | -106.3138
14.7634 | -106.3252
14.7672 | -106.3368
14.771 | -106.3503
14.7747 | -106.3639
14.7785 | -106.3748
14.7822 | -106.3846
14.786 | -106.3913
14.7898 | -106.4013
14.7935 | -106.4129
14.7973 | -106.4247
14.801 | -106.4391
14.8048 | -106.4522
14.8086 | -106.4619
14.8123 | -106.4702
14.8161 | -106.4788
14.8198 | -106.489
14.8236 | -106.4988
14.8273 | -106.508
14.8311 | -106.5127
14.8349 | -106.5185
14.8386 | -106.5218
14.8424 | -106.5254
14.8461 | -106.5287
14.8499 | -106.5289
14.8537 | -106.5275
14.8574 | -106.5228
14.8612 | -106.5174
14.8649 | -106.5141
14.8687 | -106.5083
14.8724 | -106.5008
14.8762 | -106.4926
14.88 | -106.4847
14.8837 | -106.4773
14.8875 | -106.4689
14.8912 | -106.4568
14.895 | -106.4443
14.8988 | -106.4317
14.9025 | -106.4187
14.9063 | -106.4068
14.91 | -106.3941
14.9138 | -106.3794
14.9175 | -106.3661
14.9213 | -106.3577
14.9251 | -106.3491
14.9288 | -106.34
14.9326 | -106.332
14.9363 | -106.326
14.9401 | -106.3217
14.9439 | -106.3191
14.9476 | -106.3206
14.9514 | -106.321
14.9551 | -106.3228
14.9589 | -106.3231
14.9627 | -106.324
14.9664 | -106.3269
14.9702 | -106.3315
14.9739 | -106.3399
14.9777 | -106.3471
14.9814 | -106.3498
14.9852 | -106.3426
14.989 | -106.3323
14.9927 | -106.3286
14.9965 | -106.3252
15.0002 | -106.3197
15.004 | -106.3083
15.0078 | -106.2939
15.0115 | -106.2784
15.0153 | -106.2581
15.019 | -106.2353
15.0228 | -106.2105
15.0265 | -106.1857
15.0303 | -106.1525
15.0341 | -106.1176
15.0378 | -106.0839
15.0416 | -106.0484
15.0453 | -106.0118
15.0491 | -105.9729
15.0529 | -105.934
15.0566 | -105.8957
15.0604 | -105.8607
15.0641 | -105.8274
15.0679 | -105.7943
15.0716 | -105.7654
15.0754 | -105.7374
15.0792 | -105.7134
15.0829 | -105.6931
15.0867 | -105.6748
15.0904 | -105.6625
15.0942 | -105.6545
15.098 | -105.6495
15.1017 | -105.6486
15.1055 | -105.6548
15.1092 | -105.6654
15.113 | -105.6771
15.1167 | -105.6899
15.1205 | -105.7062
15.1243 | -105.7277
15.128 | -105.7502
15.1318 | -105.7698
15.1355 | -105.7907
15.1393 | -105.8142
15.1431 | -105.8371
15.1468 | -105.8591
15.1506 | -105.8837
15.1543 | -105.9075
15.1581 | -105.9293
15.1619 | -105.9524
15.1656 | -105.9702
15.1694 | -105.985
15.1731 | -106.0024
15.1769 | -106.0162
15.1806 | -106.0206
15.1844 | -106.0308
15.1882 | -106.0417
15.1919 | -106.0512
15.1957 | -106.0602
15.1994 | -106.0718
15.2032 | -106.0884
15.207 | -106.1035
15.2107 | -106.1082
15.2145 | -106.1085
15.2182 | -106.1111
15.222 | -106.113
15.2257 | -106.1132
15.2295 | -106.1141
15.2333 | -106.1169
15.237 | -106.1192
15.2408 | -106.1212
15.2445 | -106.1238
15.2483 | -106.1288
15.2521 | -106.1307
15.2558 | -106.1302
15.2596 | -106.1289
15.2633 | -106.1283
15.2671 | -106.1259
15.2708 | -106.1255
15.2746 | -106.1232
15.2784 | -106.1166
15.2821 | -106.1113
15.2859 | -106.1099
15.2896 | -106.1051
15.2934 | -106.1002
15.2972 | -106.0936
15.3009 | -106.0858
15.3047 | -106.079
15.3084 | -106.069
15.3122 | -106.0583
15.316 | -106.0457
15.3197 | -106.0331
15.3235 | -106.0191
15.3272 | -106.0033
15.331 | -105.9879
15.3347 | -105.9717
15.3385 | -105.9579
15.3423 | -105.9424
15.346 | -105.9272
15.3498 | -105.9105
15.3535 | -105.8902
15.3573 | -105.8724
15.3611 | -105.8572
15.3648 | -105.8435
15.3686 | -105.8266
15.3723 | -105.8098
15.3761 | -105.7951
15.3798 | -105.7822
15.3836 | -105.7709
15.3874 | -105.7601
15.3911 | -105.7461
15.3949 | -105.7357
15.3986 | -105.7268
15.4024 | -105.7192
15.4062 | -105.7129
15.4099 | -105.7049
15.4137 | -105.6952
15.4174 | -105.6869
15.4212 | -105.6795
15.4249 | -105.6746
15.4287 | -105.669
15.4325 | -105.6626
15.4362 | -105.658
15.44 | -105.6526
15.4437 | -105.6449
15.4475 | -105.6398
15.4513 | -105.633
15.455 | -105.6248
15.4588 | -105.6195
15.4625 | -105.611
15.4663 | -105.602
15.47 | -105.595
15.4738 | -105.5888
15.4776 | -105.5823
15.4813 | -105.5755
15.4851 | -105.5706
15.4888 | -105.569
15.4926 | -105.5658
15.4964 | -105.5615
15.5001 | -105.56
15.5039 | -105.5556
15.5076 | -105.5471
15.5114 | -105.541
15.5152 | -105.5378
15.5189 | -105.5344
15.5227 | -105.5319
15.5264 | -105.5311
15.5302 | -105.5309
15.5339 | -105.5323
15.5377 | -105.5362
15.5415 | -105.5368
15.5452 | -105.5351
15.549 | -105.5351
15.5527 | -105.5349
15.5565 | -105.5358
15.5603 | -105.5336
15.564 | -105.5345
15.5678 | -105.533
15.5715 | -105.5327
15.5753 | -105.5375
15.579 | -105.5367
15.5828 | -105.5364
15.5866 | -105.5381
15.5903 | -105.5417
15.5941 | -105.5447
15.5978 | -105.5474
15.6016 | -105.5483
15.6054 | -105.5519
15.6091 | -105.556
15.6129 | -105.559
15.6166 | -105.5607
15.6204 | -105.5622
15.6241 | -105.5648
15.6279 | -105.5646
15.6317 | -105.5639
15.6354 | -105.565
15.6392 | -105.5677
15.6429 | -105.5699
15.6467 | -105.5698
15.6505 | -105.5709
15.6542 | -105.5734
15.658 | -105.5739
15.6617 | -105.5735
15.6655 | -105.5775
15.6693 | -105.5801
15.673 | -105.5831
15.6768 | -105.587
15.6805 | -105.5878
15.6843 | -105.5924
15.688 | -105.5961
15.6918 | -105.5963
15.6956 | -105.599
15.6993 | -105.6068
15.7031 | -105.6143
15.7068 | -105.6192
15.7106 | -105.6241
15.7144 | -105.6285
15.7181 | -105.6353
15.7219 | -105.6409
15.7256 | -105.6462
15.7294 | -105.6537
15.7331 | -105.6626
15.7369 | -105.67
15.7407 | -105.6761
15.7444 | -105.6814
15.7482 | -105.6882
15.7519 | -105.6947
15.7557 | -105.6995
15.7595 | -105.6993
15.7632 | -105.7002
15.767 | -105.707
15.7707 | -105.7171
15.7745 | -105.7246
15.7782 | -105.731
15.782 | -105.7342
15.7858 | -105.7359
15.7895 | -105.7346
15.7933 | -105.7338
15.797 | -105.7339
15.8008 | -105.7298
15.8046 | -105.7279
15.8083 | -105.7243
15.8121 | -105.7163
15.8158 | -105.7063
15.8196 | -105.6972
15.8233 | -105.687
15.8271 | -105.6774
15.8309 | -105.667
15.8346 | -105.6565
15.8384 | -105.6489
15.8421 | -105.6362
15.8459 | -105.6215
15.8497 | -105.608
15.8534 | -105.5959
15.8572 | -105.5823
15.8609 | -105.5701
15.8647 | -105.5556
15.8685 | -105.5411
15.8722 | -105.5301
15.876 | -105.5192
15.8797 | -105.5092
15.8835 | -105.498
15.8872 | -105.4876
15.891 | -105.4779
15.8948 | -105.469
15.8985 | -105.4623
15.9023 | -105.4562
15.906 | -105.453
15.9098 | -105.451
15.9136 | -105.4524
15.9173 | -105.4549
15.9211 | -105.4547
15.9248 | -105.454
15.9286 | -105.4586
15.9323 | -105.4656
15.9361 | -105.4719
15.9399 | -105.4809
15.9436 | -105.49
15.9474 | -105.5012
15.9511 | -105.5132
15.9549 | -105.5248
15.9587 | -105.5369
15.9624 | -105.5491
15.9662 | -105.5622
15.9699 | -105.5735
15.9737 | -105.5862
15.9774 | -105.5967
15.9812 | -105.6067
15.985 | -105.6187
15.9887 | -105.6325
15.9925 | -105.6426
15.9962 | -105.6479
16 | -105.6505
16.0038 | -105.6533
16.0075 | -105.6569
16.0113 | -105.6623
16.015 | -105.6683
16.0188 | -105.6739
16.0226 | -105.678
16.0263 | -105.6784
16.0301 | -105.6781
16.0338 | -105.6792
16.0376 | -105.6815
16.0413 | -105.6849
16.0451 | -105.6885
16.0489 | -105.6927
16.0526 | -105.6955
16.0564 | -105.697
16.0601 | -105.6977
16.0639 | -105.6973
16.0677 | -105.6954
16.0714 | -105.7002
16.0752 | -105.7061
16.0789 | -105.7096
16.0827 | -105.7149
16.0864 | -105.7256
16.0902 | -105.7333
16.094 | -105.7387
16.0977 | -105.7455
16.1015 | -105.7503
16.1052 | -105.7538
16.109 | -105.7588
16.1128 | -105.7658
16.1165 | -105.7735
16.1203 | -105.7785
16.124 | -105.7812
16.1278 | -105.7834
16.1315 | -105.7887
16.1353 | -105.7896
16.1391 | -105.7887
16.1428 | -105.788
16.1466 | -105.7901
16.1503 | -105.7925
16.1541 | -105.7937
16.1579 | -105.7933
16.1616 | -105.7943
16.1654 | -105.7984
16.1691 | -105.8026
16.1729 | -105.8082
16.1766 | -105.814
16.1804 | -105.8222
16.1842 | -105.8332
16.1879 | -105.8419
16.1917 | -105.8535
16.1954 | -105.8672
16.1992 | -105.881
16.203 | -105.8924
16.2067 | -105.9076
16.2105 | -105.9261
16.2142 | -105.946
16.218 | -105.9659
16.2218 | -105.9864
16.2255 | -106.0085
16.2293 | -106.0327
16.233 | -106.0561
16.2368 | -106.0764
16.2405 | -106.0963
16.2443 | -106.1157
16.2481 | -106.1318
16.2518 | -106.1492
16.2556 | -106.1663
16.2593 | -106.1803
16.2631 | -106.1896
16.2669 | -106.1827
16.2706 | -106.184
16.2744 | -106.1849
16.2781 | -106.1855
16.2819 | -106.1829
16.2856 | -106.1796
16.2894 | -106.1755
16.2932 | -106.1687
16.2969 | -106.1604
16.3007 | -106.1504
16.3044 | -106.136
16.3082 | -106.1175
16.312 | -106.0989
16.3157 | -106.0808
16.3195 | -106.0609
16.3232 | -106.0426
16.327 | -106.0275
16.3307 | -106.0107
16.3345 | -105.9939
16.3383 | -105.9799
16.342 | -105.9672
16.3458 | -105.956
16.3495 | -105.9443
16.3533 | -105.9344
16.3571 | -105.9273
16.3608 | -105.9211
16.3646 | -105.916
16.3683 | -105.916
16.3721 | -105.9161
16.3759 | -105.9163
16.3796 | -105.9158
16.3834 | -105.9197
16.3871 | -105.9253
16.3909 | -105.9326
16.3946 | -105.9405
16.3984 | -105.9471
16.4022 | -105.9538
16.4059 | -105.9622
16.4097 | -105.9717
16.4134 | -105.9806
16.4172 | -105.9904
16.421 | -106.0042
16.4247 | -106.0172
16.4285 | -106.0288
16.4322 | -106.0438
16.436 | -106.0606
16.4397 | -106.0748
16.4435 | -106.0855
16.4473 | -106.0973
16.451 | -106.1118
16.4548 | -106.1262
16.4585 | -106.1371
16.4623 | -106.1505
16.4661 | -106.1662
16.4698 | -106.1826
16.4736 | -106.1972
16.4773 | -106.212
16.4811 | -106.2285
16.4848 | -106.2444
16.4886 | -106.2616
16.4924 | -106.2796
16.4961 | -106.2997
16.4999 | -106.3212
16.5036 | -106.3422
16.5074 | -106.3637
16.5112 | -106.3843
16.5149 | -106.4058
16.5187 | -106.4271
16.5224 | -106.4467
16.5262 | -106.4654
16.53 | -106.4864
16.5337 | -106.5066
16.5375 | -106.5248
16.5412 | -106.5398
16.545 | -106.5559
16.5487 | -106.5706
16.5525 | -106.5818
16.5563 | -106.5913
16.56 | -106.5988
16.5638 | -106.6023
16.5675 | -106.6049
16.5713 | -106.6057
16.5751 | -106.6066
16.5788 | -106.6043
16.5826 | -106.599
16.5863 | -106.5914
16.5901 | -106.5813
16.5938 | -106.5712
16.5976 | -106.5613
16.6014 | -106.5507
16.6051 | -106.5383
16.6089 | -106.5266
16.6126 | -106.5148
16.6164 | -106.5028
16.6202 | -106.4881
16.6239 | -106.4741
16.6277 | -106.4617
16.6314 | -106.4509
16.6352 | -106.4407
16.6389 | -106.4339
16.6427 | -106.4262
16.6465 | -106.4196
16.6502 | -106.4143
16.654 | -106.411
16.6577 | -106.4073
16.6615 | -106.4066
16.6653 | -106.4078
16.669 | -106.4088
16.6728 | -106.4106
16.6765 | -106.4136
16.6803 | -106.42
16.684 | -106.4258
16.6878 | -106.4331
16.6916 | -106.4429
16.6953 | -106.4574
16.6991 | -106.4676
16.7028 | -106.4779
16.7066 | -106.49
16.7104 | -106.501
16.7141 | -106.5123
16.7179 | -106.5253
16.7216 | -106.5369
16.7254 | -106.549
16.7292 | -106.5595
16.7329 | -106.568
16.7367 | -106.5763
16.7404 | -106.5857
16.7442 | -106.5943
16.7479 | -106.601
16.7517 | -106.6062
16.7555 | -106.6102
16.7592 | -106.6155
16.763 | -106.6181
16.7667 | -106.6184
16.7705 | -106.6185
16.7743 | -106.6183
16.778 | -106.6188
16.7818 | -106.6184
16.7855 | -106.6166
16.7893 | -106.6139
16.793 | -106.6087
16.7968 | -106.6035
16.8006 | -106.6
16.8043 | -106.5982
16.8081 | -106.5997
16.8118 | -106.5983
16.8156 | -106.5955
16.8194 | -106.5914
16.8231 | -106.5863
16.8269 | -106.585
16.8306 | -106.5851
16.8344 | -106.5835
16.8381 | -106.5815
16.8419 | -106.5786
16.8457 | -106.5783
16.8494 | -106.5785
16.8532 | -106.5776
16.8569 | -106.5769
16.8607 | -106.5773
16.8645 | -106.5768
16.8682 | -106.5758
16.872 | -106.5743
16.8757 | -106.5756
16.8795 | -106.5762
16.8833 | -106.5765
16.887 | -106.5772
16.8908 | -106.5781
16.8945 | -106.5779
16.8983 | -106.5799
16.902 | -106.5804
16.9058 | -106.581
16.9096 | -106.5832
16.9133 | -106.589
16.9171 | -106.595
16.9208 | -106.5991
16.9246 | -106.6051
16.9284 | -106.6118
16.9321 | -106.6195
16.9359 | -106.6282
16.9396 | -106.64
16.9434 | -106.653
16.9471 | -106.6654
16.9509 | -106.6792
16.9547 | -106.6945
16.9584 | -106.71
16.9622 | -106.7273
16.9659 | -106.7417
16.9697 | -106.7579
16.9735 | -106.7753
16.9772 | -106.7925
16.981 | -106.8099
16.9847 | -106.8291
16.9885 | -106.8472
16.9922 | -106.8635
16.996 | -106.881
16.9998 | -106.8989
17.0035 | -106.917
17.0073 | -106.9301
17.011 | -106.941
17.0148 | -106.9536
17.0186 | -106.9683
17.0223 | -106.9783
17.0261 | -106.9879
17.0298 | -106.999
17.0336 | -107.0101
17.0373 | -107.0193
17.0411 | -107.0267
17.0449 | -107.0303
17.0486 | -107.0316
17.0524 | -107.0324
17.0561 | -107.034
17.0599 | -107.0351
17.0637 | -107.0371
17.0674 | -107.0397
17.0712 | -107.0386
17.0749 | -107.0349
17.0787 | -107.0283
17.0825 | -107.0203
17.0862 | -107.0141
17.09 | -107.0086
17.0937 | -107.0046
17.0975 | -107.0003
17.1012 | -106.9973
17.105 | -106.9931
17.1088 | -106.9895
17.1125 | -106.9872
17.1163 | -106.9862
17.12 | -106.9833
17.1238 | -106.9794
17.1276 | -106.9786
17.1313 | -106.9765
17.1351 | -106.977
17.1388 | -106.9775
17.1426 | -106.9791
17.1463 | -106.9838
17.1501 | -106.988
17.1539 | -106.9919
17.1576 | -106.9924
17.1614 | -106.9946
17.1651 | -106.9995
17.1689 | -107.0018
17.1727 | -107.0059
17.1764 | -107.0113
17.1802 | -107.014
17.1839 | -107.0182
17.1877 | -107.0223
17.1914 | -107.0276
17.1952 | -107.0336
17.199 | -107.0362
17.2027 | -107.0413
17.2065 | -107.0456
17.2102 | -107.0504
17.214 | -107.0559
17.2178 | -107.0613
17.2215 | -107.0647
17.2253 | -107.0686
17.229 | -107.0708
17.2328 | -107.0762
17.2366 | -107.0809
17.2403 | -107.0865
17.2441 | -107.0897
17.2478 | -107.093
17.2516 | -107.0986
17.2553 | -107.1025
17.2591 | -107.1047
17.2629 | -107.105
17.2666 | -107.1087
17.2704 | -107.1132
17.2741 | -107.1166
17.2779 | -107.1196
17.2817 | -107.1244
17.2854 | -107.1302
17.2892 | -107.1364
17.2929 | -107.1426
17.2967 | -107.1493
17.3004 | -107.1572
17.3042 | -107.1648
17.308 | -107.1736
17.3117 | -107.1827
17.3155 | -107.1939
17.3192 | -107.2054
17.323 | -107.2182
17.3268 | -107.2311
17.3305 | -107.2421
17.3343 | -107.2537
17.338 | -107.2644
17.3418 | -107.2762
17.3455 | -107.287
17.3493 | -107.2919
17.3531 | -107.3005
17.3568 | -107.3106
17.3606 | -107.3169
17.3643 | -107.3212
17.3681 | -107.3283
17.3719 | -107.3362
17.3756 | -107.34
17.3794 | -107.3434
17.3831 | -107.3425
17.3869 | -107.3436
17.3907 | -107.3434
17.3944 | -107.3389
17.3982 | -107.3361
17.4019 | -107.3316
17.4057 | -107.3275
17.4094 | -107.3222
17.4132 | -107.3134
17.417 | -107.3061
17.4207 | -107.2957
17.4245 | -107.2844
17.4282 | -107.2758
17.432 | -107.2662
17.4358 | -107.2554
17.4395 | -107.2438
17.4433 | -107.2332
17.447 | -107.2243
17.4508 | -107.2083
17.4545 | -107.196
17.4583 | -107.1874
17.4621 | -107.1795
17.4658 | -107.1704
17.4696 | -107.1632
17.4733 | -107.155
17.4771 | -107.1504
17.4809 | -107.1444
17.4846 | -107.1387
17.4884 | -107.1319
17.4921 | -107.126
17.4959 | -107.1209
17.4996 | -107.1197
17.5034 | -107.1209
17.5072 | -107.1201
17.5109 | -107.1188
17.5147 | -107.117
17.5184 | -107.1139
17.5222 | -107.1111
17.526 | -107.1093
17.5297 | -107.1069
17.5335 | -107.1054
17.5372 | -107.1045
17.541 | -107.1049
17.5448 | -107.1062
17.5485 | -107.106
17.5523 | -107.1052
17.556 | -107.103
17.5598 | -107.1005
17.5635 | -107.0983
17.5673 | -107.0958
17.5711 | -107.0934
17.5748 | -107.0903
17.5786 | -107.0847
17.5823 | -107.077
17.5861 | -107.0703
17.5899 | -107.0652
17.5936 | -107.0592
17.5974 | -107.0541
17.6011 | -107.051
17.6049 | -107.0443
17.6086 | -107.0369
17.6124 | -107.031
17.6162 | -107.0253
17.6199 | -107.0181
17.6237 | -107.0094
17.6274 | -107.001
17.6312 | -106.9919
17.635 | -106.9839
17.6387 | -106.9755
17.6425 | -106.9668
17.6462 | -106.9585
17.65 | -106.9472
17.6537 | -106.9367
17.6575 | -106.9278
17.6613 | -106.9189
17.665 | -106.9094
17.6688 | -106.8988
17.6725 | -106.889
17.6763 | -106.8783
17.6801 | -106.8671
17.6838 | -106.8571
17.6876 | -106.8476
17.6913 | -106.8374
17.6951 | -106.8289
17.6988 | -106.8204
17.7026 | -106.8156
17.7064 | -106.8102
17.7101 | -106.8031
17.7139 | -106.7926
17.7176 | -106.7832
17.7214 | -106.778
17.7252 | -106.7739
17.7289 | -106.7693
17.7327 | -106.7649
17.7364 | -106.7634
17.7402 | -106.7652
17.744 | -106.7681
17.7477 | -106.7725
17.7515 | -106.7746
17.7552 | -106.7775
17.759 | -106.7831
17.7627 | -106.7893
17.7665 | -106.7959
17.7703 | -106.8045
17.774 | -106.8133
17.7778 | -106.8232
17.7815 | -106.8319
17.7853 | -106.8423
17.7891 | -106.8531
17.7928 | -106.8629
17.7966 | -106.874
17.8003 | -106.8843
17.8041 | -106.8955
17.8078 | -106.9102
17.8116 | -106.9259
17.8154 | -106.9386
17.8191 | -106.9493
17.8229 | -106.9614
17.8266 | -106.9742
17.8304 | -106.9902
17.8342 | -107.0046
17.8379 | -107.015
17.8417 | -107.0253
17.8454 | -107.0371
17.8492 | -107.0466
17.8529 | -107.0569
17.8567 | -107.0657
17.8605 | -107.0702
17.8642 | -107.0744
17.868 | -107.081
17.8717 | -107.0868
17.8755 | -107.0918
17.8793 | -107.0977
17.883 | -107.1027
17.8868 | -107.1035
17.8905 | -107.0971
17.8943 | -107.0946
17.8981 | -107.0919
17.9018 | -107.0923
17.9056 | -107.0912
17.9093 | -107.0859
17.9131 | -107.08
17.9168 | -107.0717
17.9206 | -107.0619
17.9244 | -107.0543
17.9281 | -107.0458
17.9319 | -107.0376
17.9356 | -107.0266
17.9394 | -107.0135
17.9432 | -107
17.9469 | -106.9873
17.9507 | -106.9719
17.9544 | -106.9557
17.9582 | -106.9401
17.9619 | -106.9226
17.9657 | -106.9063
17.9695 | -106.8918
17.9732 | -106.876
17.977 | -106.8577
17.9807 | -106.8383
17.9845 | -106.819
17.9883 | -106.8015
17.992 | -106.7814
17.9958 | -106.7602
17.9995 | -106.7394
18.0033 | -106.7192
18.007 | -106.7011
18.0108 | -106.6836
18.0146 | -106.6684
18.0183 | -106.6542
18.0221 | -106.6421
18.0258 | -106.6322
18.0296 | -106.6236
18.0334 | -106.6161
18.0371 | -106.613
18.0409 | -106.6081
18.0446 | -106.6077
18.0484 | -106.6098
18.0521 | -106.6118
18.0559 | -106.6182
18.0597 | -106.6273
18.0634 | -106.6358
18.0672 | -106.645
18.0709 | -106.6582
18.0747 | -106.675
18.0785 | -106.6909
18.0822 | -106.7092
18.086 | -106.7264
18.0897 | -106.744
18.0935 | -106.7647
18.0973 | -106.7842
18.101 | -106.8007
18.1048 | -106.8181
18.1085 | -106.835
18.1123 | -106.8526
18.116 | -106.8671
18.1198 | -106.88
18.1236 | -106.8916
18.1273 | -106.9
18.1311 | -106.9068
18.1348 | -106.9111
18.1386 | -106.9141
18.1424 | -106.9184
18.1461 | -106.9219
18.1499 | -106.9204
18.1536 | -106.9162
18.1574 | -106.9121
18.1611 | -106.9043
18.1649 | -106.8983
18.1687 | -106.8936
18.1724 | -106.8874
18.1762 | -106.882
18.1799 | -106.8775
18.1837 | -106.8691
18.1875 | -106.8616
18.1912 | -106.8526
18.195 | -106.844
18.1987 | -106.8378
18.2025 | -106.8339
18.2062 | -106.8309
18.21 | -106.8278
18.2138 | -106.8259
18.2175 | -106.8228
18.2213 | -106.8224
18.225 | -106.8208
18.2288 | -106.8194
18.2326 | -106.8207
18.2363 | -106.8232
18.2401 | -106.8261
18.2438 | -106.8306
18.2476 | -106.836
18.2514 | -106.8415
18.2551 | -106.8468
18.2589 | -106.8512
18.2626 | -106.8544
18.2664 | -106.8578
18.2701 | -106.8611
18.2739 | -106.8621
18.2777 | -106.8623
18.2814 | -106.8628
18.2852 | -106.861
18.2889 | -106.8584
18.2927 | -106.8558
18.2965 | -106.8525
18.3002 | -106.8481
18.304 | -106.8411
18.3077 | -106.8309
18.3115 | -106.8196
18.3152 | -106.8058
18.319 | -106.79
18.3228 | -106.7723
18.3265 | -106.7552
18.3303 | -106.7363
18.334 | -106.7174
18.3378 | -106.6971
18.3416 | -106.6749
18.3453 | -106.6523
18.3491 | -106.631
18.3528 | -106.6075
18.3566 | -106.5834
18.3603 | -106.5619
18.3641 | -106.541
18.3679 | -106.519
18.3716 | -106.4981
18.3754 | -106.4822
18.3791 | -106.4684
18.3829 | -106.4544
18.3867 | -106.4427
18.3904 | -106.4343
18.3942 | -106.4284
18.3979 | -106.4216
18.4017 | -106.4152
18.4054 | -106.4121
18.4092 | -106.412
18.413 | -106.4154
18.4167 | -106.4205
18.4205 | -106.4268
18.4242 | -106.4356
18.428 | -106.4451
18.4318 | -106.4554
18.4355 | -106.476
18.4393 | -106.4901
18.443 | -106.5047
18.4468 | -106.5201
18.4506 | -106.5356
18.4543 | -106.5534
18.4581 | -106.5706
18.4618 | -106.586
18.4656 | -106.6
18.4693 | -106.6147
18.4731 | -106.6284
18.4769 | -106.643
18.4806 | -106.6581
18.4844 | -106.6701
18.4881 | -106.6834
18.4919 | -106.6974
18.4957 | -106.7086
18.4994 | -106.7186
18.5032 | -106.7301
18.5069 | -106.7391
18.5107 | -106.7471
18.5144 | -106.756
18.5182 | -106.7658
18.522 | -106.7773
18.5257 | -106.7864
18.5295 | -106.7952
18.5332 | -106.8046
18.537 | -106.8143
18.5408 | -106.8243
18.5445 | -106.8317
18.5483 | -106.8388
18.552 | -106.8456
18.5558 | -106.8518
18.5595 | -106.8585
18.5633 | -106.8647
18.5671 | -106.8701
18.5708 | -106.8774
18.5746 | -106.8852
18.5783 | -106.8938
18.5821 | -106.9007
18.5859 | -106.908
18.5896 | -106.9147
18.5934 | -106.9207
18.5971 | -106.9244
18.6009 | -106.93
18.6047 | -106.9368
18.6084 | -106.9438
18.6122 | -106.9517
18.6159 | -106.9613
18.6197 | -106.9712
18.6234 | -106.9804
18.6272 | -106.9898
18.631 | -106.999
18.6347 | -107.0093
18.6385 | -107.0188
18.6422 | -107.0267
18.646 | -107.0334
18.6498 | -107.0425
18.6535 | -107.0518
18.6573 | -107.0621
18.661 | -107.0774
18.6648 | -107.0911
18.6685 | -107.1022
18.6723 | -107.115
18.6761 | -107.1246
18.6798 | -107.1338
18.6836 | -107.1429
18.6873 | -107.1512
18.6911 | -107.1608
18.6949 | -107.1691
18.6986 | -107.177
18.7024 | -107.1835
18.7061 | -107.1904
18.7099 | -107.199
18.7136 | -107.2073
18.7174 | -107.2141
18.7212 | -107.2205
18.7249 | -107.2303
18.7287 | -107.2386
18.7324 | -107.2459
18.7362 | -107.25
18.74 | -107.2555
18.7437 | -107.2598
18.7475 | -107.2642
18.7512 | -107.269
18.755 | -107.2728
18.7587 | -107.2742
18.7625 | -107.2792
18.7663 | -107.2836
18.77 | -107.2869
18.7738 | -107.2895
18.7775 | -107.2908
18.7813 | -107.2929
18.7851 | -107.2935
18.7888 | -107.2951
18.7926 | -107.2981
18.7963 | -107.2987
18.8001 | -107.2997
18.8039 | -107.3012
18.8076 | -107.3027
18.8114 | -107.3058
18.8151 | -107.308
18.8189 | -107.3103
18.8226 | -107.3117
18.8264 | -107.3107
18.8302 | -107.312
18.8339 | -107.3113
18.8377 | -107.3089
18.8414 | -107.3066
18.8452 | -107.3036
18.849 | -107.3012
18.8527 | -107.295
18.8565 | -107.2881
18.8602 | -107.2824
18.864 | -107.2776
18.8677 | -107.2727
18.8715 | -107.2642
18.8753 | -107.2559
18.879 | -107.2458
18.8828 | -107.2364
18.8865 | -107.2252
18.8903 | -107.2149
18.8941 | -107.2042
18.8978 | -107.1917
18.9016 | -107.1786
18.9053 | -107.1664
18.9091 | -107.1537
18.9128 | -107.1431
18.9166 | -107.1302
18.9204 | -107.1176
18.9241 | -107.1042
18.9279 | -107.0913
18.9316 | -107.0802
18.9354 | -107.0704
18.9392 | -107.06
18.9429 | -107.0503
18.9467 | -107.0409
18.9504 | -107.0339
18.9542 | -107.0261
18.958 | -107.0187
18.9617 | -107.0122
18.9655 | -107.0066
18.9692 | -107.0043
18.973 | -106.9999
18.9767 | -107.0042
18.9805 | -107.0029
18.9843 | -107.0039
18.988 | -107.0069
18.9918 | -107.0097
18.9955 | -107.0129
18.9993 | -107.0186
19.0031 | -107.0264
19.0068 | -107.0375
19.0106 | -107.0489
19.0143 | -107.059
19.0181 | -107.071
19.0218 | -107.0841
19.0256 | -107.096
19.0294 | -107.1098
19.0331 | -107.1231
19.0369 | -107.1361
19.0406 | -107.1487
19.0444 | -107.1622
19.0482 | -107.1777
19.0519 | -107.1913
19.0557 | -107.204
19.0594 | -107.2138
19.0632 | -107.2246
19.0669 | -107.2365
19.0707 | -107.2456
19.0745 | -107.2545
19.0782 | -107.2618
19.082 | -107.2678
19.0857 | -107.2731
19.0895 | -107.2772
19.0933 | -107.281
19.097 | -107.2823
19.1008 | -107.2826
19.1045 | -107.2831
19.1083 | -107.2837
19.1121 | -107.2827
19.1158 | -107.2803
19.1196 | -107.2767
19.1233 | -107.2744
19.1271 | -107.2708
19.1308 | -107.2657
19.1346 | -107.2593
19.1384 | -107.2511
19.1421 | -107.245
19.1459 | -107.2375
19.1496 | -107.2317
19.1534 | -107.2237
19.1572 | -107.2168
19.1609 | -107.2097
19.1647 | -107.2023
19.1684 | -107.1974
19.1722 | -107.1931
19.1759 | -107.1847
19.1797 | -107.1774
19.1835 | -107.1701
19.1872 | -107.165
19.191 | -107.1591
19.1947 | -107.1537
19.1985 | -107.1477
19.2023 | -107.1429
19.206 | -107.1383
19.2098 | -107.1342
19.2135 | -107.1317
19.2173 | -107.1286
19.221 | -107.1236
19.2248 | -107.1159
19.2286 | -107.1097
19.2323 | -107.1069
19.2361 | -107.1037
19.2398 | -107.0984
19.2436 | -107.0939
19.2474 | -107.086
19.2511 | -107.0798
19.2549 | -107.0731
19.2586 | -107.0667
19.2624 | -107.0579
19.2661 | -107.0482
19.2699 | -107.038
19.2737 | -107.0282
19.2774 | -107.0188
19.2812 | -107.0062
19.2849 | -106.993
19.2887 | -106.9806
19.2925 | -106.9685
19.2962 | -106.9542
19.3 | -106.94
19.3037 | -106.9253
19.3075 | -106.9101
19.3113 | -106.8939
19.315 | -106.8783
19.3188 | -106.8623
19.3225 | -106.8449
19.3263 | -106.8288
19.33 | -106.8123
19.3338 | -106.7971
19.3376 | -106.7824
19.3413 | -106.7669
19.3451 | -106.7524
19.3488 | -106.7379
19.3526 | -106.7258
19.3564 | -106.7151
19.3601 | -106.7068
19.3639 | -106.698
19.3676 | -106.6897
19.3714 | -106.6835
19.3751 | -106.6789
19.3789 | -106.6744
19.3827 | -106.6713
19.3864 | -106.6663
19.3902 | -106.6638
19.3939 | -106.6644
19.3977 | -106.6651
19.4015 | -106.6677
19.4052 | -106.6718
19.409 | -106.677
19.4127 | -106.683
19.4165 | -106.6884
19.4202 | -106.6936
19.424 | -106.6999
19.4278 | -106.709
19.4315 | -106.7176
19.4353 | -106.7253
19.439 | -106.7328
19.4428 | -106.7413
19.4466 | -106.7471
19.4503 | -106.7566
19.4541 | -106.7646
19.4578 | -106.7722
19.4616 | -106.7784
19.4654 | -106.7873
19.4691 | -106.7965
19.4729 | -106.8035
19.4766 | -106.8102
19.4804 | -106.8178
19.4841 | -106.8247
19.4879 | -106.8311
19.4917 | -106.8363
19.4954 | -106.8415
19.4992 | -106.8493
19.5029 | -106.8562
19.5067 | -106.8608
19.5105 | -106.8646
19.5142 | -106.8695
19.518 | -106.8751
19.5217 | -106.8797
19.5255 | -106.8865
19.5292 | -106.8916
19.533 | -106.8971
19.5368 | -106.9026
19.5405 | -106.9095
19.5443 | -106.9166
19.548 | -106.9239
19.5518 | -106.9299
19.5556 | -106.9353
19.5593 | -106.9426
19.5631 | -106.9511
19.5668 | -106.959
19.5706 | -106.9667
19.5743 | -106.9768
19.5781 | -106.9856
19.5819 | -106.9938
19.5856 | -107.0019
19.5894 | -107.0106
19.5931 | -107.0209
19.5969 | -107.0311
19.6007 | -107.0393
19.6044 | -107.047
19.6082 | -107.0556
19.6119 | -107.0622
19.6157 | -107.0684
19.6195 | -107.076
19.6232 | -107.0831
19.627 | -107.0898
19.6307 | -107.0955
19.6345 | -107.0988
19.6382 | -107.1018
19.642 | -107.1077
19.6458 | -107.1141
19.6495 | -107.1195
19.6533 | -107.1229
19.657 | -107.1268
19.6608 | -107.1311
19.6646 | -107.1369
19.6683 | -107.1405
19.6721 | -107.1434
19.6758 | -107.1477
19.6796 | -107.1488
19.6833 | -107.1474
19.6871 | -107.1451
19.6909 | -107.1439
19.6946 | -107.1447
19.6984 | -107.1448
19.7021 | -107.1416
19.7059 | -107.1356
19.7097 | -107.129
19.7134 | -107.1245
19.7172 | -107.1206
19.7209 | -107.1125
19.7247 | -107.1047
19.7284 | -107.097
19.7322 | -107.0886
19.736 | -107.0768
19.7397 | -107.0644
19.7435 | -107.051
19.7472 | -107.0375
19.751 | -107.0249
19.7548 | -107.0108
19.7585 | -106.9943
19.7623 | -106.9793
19.766 | -106.9644
19.7698 | -106.9492
19.7735 | -106.9326
19.7773 | -106.9157
19.7811 | -106.8948
19.7848 | -106.8723
19.7886 | -106.8534
19.7923 | -106.8345
19.7961 | -106.8163
19.7999 | -106.797
19.8036 | -106.7776
19.8074 | -106.7595
19.8111 | -106.7383
19.8149 | -106.7178
19.8187 | -106.6974
19.8224 | -106.6781
19.8262 | -106.6593
19.8299 | -106.6428
19.8337 | -106.6276
19.8374 | -106.6119
19.8412 | -106.5967
19.845 | -106.5831
19.8487 | -106.5715
19.8525 | -106.5608
19.8562 | -106.5541
19.86 | -106.5489
19.8638 | -106.5454
19.8675 | -106.5423
19.8713 | -106.5402
19.875 | -106.5386
19.8788 | -106.5412
19.8825 | -106.5477
19.8863 | -106.5551
19.8901 | -106.5641
19.8938 | -106.5759
19.8976 | -106.5869
19.9013 | -106.5979
19.9051 | -106.6112
19.9089 | -106.6263
19.9126 | -106.6432
19.9164 | -106.6624
19.9201 | -106.6814
19.9239 | -106.6997
19.9276 | -106.7181
19.9314 | -106.7388
19.9352 | -106.7599
19.9389 | -106.7814
19.9427 | -106.8012
19.9464 | -106.8199
19.9502 | -106.8394
19.954 | -106.8576
19.9577 | -106.8755
19.9615 | -106.8935
19.9652 | -106.9105
19.969 | -106.929
19.9728 | -106.9451
19.9765 | -106.9587
19.9803 | -106.9737
19.984 | -106.988
19.9878 | -107.0011
19.9915 | -107.0122
19.9953 | -107.0191
19.9991 | -107.0251
20.0028 | -107.0332
20.0066 | -107.0417
20.0103 | -107.0463
20.0141 | -107.0502
20.0179 | -107.0549
20.0216 | -107.0594
20.0254 | -107.0614
20.0291 | -107.0615
20.0329 | -107.0616
20.0366 | -107.06
20.0404 | -107.06
20.0442 | -107.0606
20.0479 | -107.0598
20.0517 | -107.0596
20.0554 | -107.0595
20.0592 | -107.0563
20.063 | -107.0528
20.0667 | -107.0513
20.0705 | -107.0521
20.0742 | -107.0492
20.078 | -107.0455
20.0817 | -107.0402
20.0855 | -107.0351
20.0893 | -107.0319
20.093 | -107.0281
20.0968 | -107.0247
20.1005 | -107.0201
20.1043 | -107.0158
20.1081 | -107.0121
20.1118 | -107.0092
20.1156 | -107.0039
20.1193 | -106.9967
20.1231 | -106.9924
20.1269 | -106.9898
20.1306 | -106.9841
20.1344 | -106.9776
20.1381 | -106.9689
20.1419 | -106.9545
20.1456 | -106.9433
20.1494 | -106.9351
20.1532 | -106.926
20.1569 | -106.9174
20.1607 | -106.9082
20.1644 | -106.8944
20.1682 | -106.8815
20.172 | -106.8691
20.1757 | -106.8555
20.1795 | -106.8421
20.1832 | -106.8292
20.187 | -106.8172
20.1907 | -106.8045
20.1945 | -106.7904
20.1983 | -106.7778
20.202 | -106.765
20.2058 | -106.7519
20.2095 | -106.7388
20.2133 | -106.7265
20.2171 | -106.7153
20.2208 | -106.7056
20.2246 | -106.6948
20.2283 | -106.6847
20.2321 | -106.6759
20.2358 | -106.6656
20.2396 | -106.6554
20.2434 | -106.6465
20.2471 | -106.6404
20.2509 | -106.6371
20.2546 | -106.6323
20.2584 | -106.6282
20.2622 | -106.6241
20.2659 | -106.6211
20.2697 | -106.6189
20.2734 | -106.618
20.2772 | -106.6193
20.2809 | -106.6222
20.2847 | -106.6243
20.2885 | -106.6261
20.2922 | -106.6278
20.296 | -106.63
20.2997 | -106.6322
20.3035 | -106.6353
20.3073 | -106.6391
20.311 | -106.6431
20.3148 | -106.6489
20.3185 | -106.6528
20.3223 | -106.6579
20.3261 | -106.6628
20.3298 | -106.6658
20.3336 | -106.6684
20.3373 | -106.6698
20.3411 | -106.6733
20.3448 | -106.6789
20.3486 | -106.6837
20.3524 | -106.6884
20.3561 | -106.6937
20.3599 | -106.6957
20.3636 | -106.6977
20.3674 | -106.6998
20.3712 | -106.7037
20.3749 | -106.7075
20.3787 | -106.7102
20.3824 | -106.7123
20.3862 | -106.7147
20.3899 | -106.7166
20.3937 | -106.7182
20.3975 | -106.7188
20.4012 | -106.718
20.405 | -106.7167
20.4087 | -106.7191
20.4125 | -106.7205
20.4163 | -106.7196
20.42 | -106.7209
20.4238 | -106.7228
20.4275 | -106.7251
20.4313 | -106.7271
20.435 | -106.7281
20.4388 | -106.7254
20.4426 | -106.7252
20.4463 | -106.7251
20.4501 | -106.7234
20.4538 | -106.7236
20.4576 | -106.7231
20.4614 | -106.7223
20.4651 | -106.7231
20.4689 | -106.7221
20.4726 | -106.7197
20.4764 | -106.7154
20.4801 | -106.7129
20.4839 | -106.7095
20.4877 | -106.7044
20.4914 | -106.7
20.4952 | -106.6928
20.4989 | -106.6875
20.5027 | -106.6826
20.5065 | -106.6778
20.5102 | -106.6728
20.514 | -106.6682
20.5177 | -106.6627
20.5215 | -106.6584
20.5253 | -106.6558
20.529 | -106.6521
20.5328 | -106.6465
20.5365 | -106.6406
20.5403 | -106.6324
20.544 | -106.6248
20.5478 | -106.6192
20.5516 | -106.6147
20.5553 | -106.6093
20.5591 | -106.6042
20.5628 | -106.6001
20.5666 | -106.5963
20.5704 | -106.592
20.5741 | -106.5891
20.5779 | -106.5865
20.5816 | -106.5848
20.5854 | -106.5844
20.5891 | -106.5852
20.5929 | -106.5856
20.5967 | -106.5846
20.6004 | -106.5833
20.6042 | -106.5894
20.6079 | -106.5892
20.6117 | -106.5909
20.6155 | -106.5934
20.6192 | -106.5954
20.623 | -106.5982
20.6267 | -106.6023
20.6305 | -106.6092
20.6342 | -106.6138
20.638 | -106.6157
20.6418 | -106.6176
20.6455 | -106.6213
20.6493 | -106.6228
20.653 | -106.6261
20.6568 | -106.6273
20.6606 | -106.6265
20.6643 | -106.6255
20.6681 | -106.6247
20.6718 | -106.6245
20.6756 | -106.6249
20.6794 | -106.624
20.6831 | -106.6188
20.6869 | -106.6128
20.6906 | -106.608
20.6944 | -106.6018
20.6981 | -106.5938
20.7019 | -106.583
20.7057 | -106.5698
20.7094 | -106.5565
20.7132 | -106.5436
20.7169 | -106.5276
20.7207 | -106.5132
20.7245 | -106.5004
20.7282 | -106.4841
20.732 | -106.4674
20.7357 | -106.4489
20.7395 | -106.4294
20.7432 | -106.4113
20.747 | -106.3923
20.7508 | -106.3718
20.7545 | -106.3521
20.7583 | -106.3328
20.762 | -106.3143
20.7658 | -106.2981
20.7696 | -106.2797
20.7733 | -106.2576
20.7771 | -106.2387
20.7808 | -106.2233
20.7846 | -106.208
20.7883 | -106.1964
20.7921 | -106.1833
20.7959 | -106.1701
20.7996 | -106.1556
20.8034 | -106.1419
20.8071 | -106.1318
20.8109 | -106.1238
20.8147 | -106.1148
20.8184 | -106.1053
20.8222 | -106.0982
20.8259 | -106.0939
20.8297 | -106.0899
20.8335 | -106.0893
20.8372 | -106.0888
20.841 | -106.0876
20.8447 | -106.0882
20.8485 | -106.0897
20.8522 | -106.0905
20.856 | -106.0923
20.8598 | -106.0973
20.8635 | -106.1046
20.8673 | -106.1117
20.871 | -106.1193
20.8748 | -106.1268
20.8786 | -106.1351
20.8823 | -106.1442
20.8861 | -106.1535
20.8898 | -106.1651
20.8936 | -106.1777
20.8973 | -106.1896
20.9011 | -106.2001
20.9049 | -106.213
20.9086 | -106.2259
20.9124 | -106.2392
20.9161 | -106.2518
20.9199 | -106.2627
20.9237 | -106.2755
20.9274 | -106.2882
20.9312 | -106.3018
20.9349 | -106.3164
20.9387 | -106.3305
20.9424 | -106.3447
20.9462 | -106.3601
20.95 | -106.3753
20.9537 | -106.3922
20.9575 | -106.4093
20.9612 | -106.4253
20.965 | -106.4421
20.9688 | -106.457
20.9725 | -106.4698
20.9763 | -106.485
20.98 | -106.5027
20.9838 | -106.5175
20.9875 | -106.5266
20.9913 | -106.5356
20.9951 | -106.5476
20.9988 | -106.5589
21.0026 | -106.5722
21.0063 | -106.584
21.0101 | -106.5925
21.0139 | -106.6012
21.0176 | -106.6082
21.0214 | -106.6151
21.0251 | -106.6224
21.0289 | -106.6272
21.0327 | -106.6331
21.0364 | -106.6402
21.0402 | -106.6454
21.0439 | -106.6525
21.0477 | -106.6574
21.0514 | -106.6576
21.0552 | -106.658
21.059 | -106.6603
21.0627 | -106.6621
21.0665 | -106.6665
21.0702 | -106.6703
21.074 | -106.6719
21.0778 | -106.6711
21.0815 | -106.6702
21.0853 | -106.6687
21.089 | -106.668
21.0928 | -106.6662
21.0965 | -106.6652
21.1003 | -106.6618
21.1041 | -106.6556
21.1078 | -106.6508
21.1116 | -106.6489
21.1153 | -106.648
21.1191 | -106.6465
21.1229 | -106.6443
21.1266 | -106.6411
21.1304 | -106.6387
21.1341 | -106.6377
21.1379 | -106.638
21.1416 | -106.6353
21.1454 | -106.6343
21.1492 | -106.6337
21.1529 | -106.6356
21.1567 | -106.636
21.1604 | -106.6337
21.1642 | -106.6303
21.168 | -106.6294
21.1717 | -106.6294
21.1755 | -106.6299
21.1792 | -106.6293
21.183 | -106.6289
21.1868 | -106.631
21.1905 | -106.6329
21.1943 | -106.6349
21.198 | -106.6365
21.2018 | -106.6393
21.2055 | -106.6414
21.2093 | -106.6446
21.2131 | -106.6477
21.2168 | -106.6507
21.2206 | -106.6542
21.2243 | -106.6552
21.2281 | -106.6548
21.2319 | -106.6566
21.2356 | -106.6602
21.2394 | -106.664
21.2431 | -106.6669
21.2469 | -106.6689
21.2506 | -106.6711
21.2544 | -106.6729
21.2582 | -106.675
21.2619 | -106.6778
21.2657 | -106.6803
21.2694 | -106.6826
21.2732 | -106.6873
21.277 | -106.6922
21.2807 | -106.6965
21.2845 | -106.7032
21.2882 | -106.7113
21.292 | -106.718
21.2957 | -106.7228
21.2995 | -106.7272
21.3033 | -106.7301
21.307 | -106.734
21.3108 | -106.7402
21.3145 | -106.7476
21.3183 | -106.7543
21.3221 | -106.7596
21.3258 | -106.7652
21.3296 | -106.7731
21.3333 | -106.7784
21.3371 | -106.7807
21.3408 | -106.785
21.3446 | -106.7872
21.3484 | -106.7918
21.3521 | -106.799
21.3559 | -106.8026
21.3596 | -106.8031
21.3634 | -106.8014
21.3672 | -106.802
21.3709 | -106.8049
21.3747 | -106.804
21.3784 | -106.7974
21.3822 | -106.7918
21.386 | -106.7853
21.3897 | -106.774
21.3935 | -106.7626
21.3972 | -106.7464
21.401 | -106.7255
21.4047 | -106.7001
21.4085 | -106.6684
21.4123 | -106.6285
21.416 | -106.5858
21.4198 | -106.5392
21.4235 | -106.4871
21.4273 | -106.4366
21.4311 | -106.3818
21.4348 | -106.3275
21.4386 | -106.2682
21.4423 | -106.2012
21.4461 | -106.1341
21.4498 | -106.0585
21.4536 | -105.9785
21.4574 | -105.8999
21.4611 | -105.8163
21.4649 | -105.7323
21.4686 | -105.6481
21.4724 | -105.5597
21.4762 | -105.4694
21.4799 | -105.379
21.4837 | -105.2892
21.4874 | -105.1985
21.4912 | -105.1048
21.4949 | -105.0118
21.4987 | -104.9209
21.5025 | -104.8305
21.5062 | -104.7419
21.51 | -104.6564
21.5137 | -104.5735
21.5175 | -104.4933
21.5213 | -104.4162
21.525 | -104.3428
21.5288 | -104.2738
21.5325 | -104.2098
21.5363 | -104.15
21.5401 | -104.0938
21.5438 | -104.0424
21.5476 | -103.9959
21.5513 | -103.9565
21.5551 | -103.9219
21.5588 | -103.8923
21.5626 | -103.8698
21.5664 | -103.8532
21.5701 | -103.8419
21.5739 | -103.8375
21.5776 | -103.8386
21.5814 | -103.8451
21.5852 | -103.8579
21.5889 | -103.8766
21.5927 | -103.902
21.5964 | -103.934
21.6002 | -103.9698
21.6039 | -104.0103
21.6077 | -104.056
21.6115 | -104.1076
21.6152 | -104.1629
21.619 | -104.2212
21.6227 | -104.2837
21.6265 | -104.3497
21.6303 | -104.4188
21.634 | -104.4913
21.6378 | -104.5681
21.6415 | -104.6486
21.6453 | -104.7301
21.649 | -104.8129
21.6528 | -104.8954
21.6566 | -104.9787
21.6603 | -105.0638
21.6641 | -105.1499
21.6678 | -105.2359
21.6716 | -105.3187
21.6754 | -105.3996
21.6791 | -105.4806
21.6829 | -105.5631
21.6866 | -105.6116
21.6904 | -105.6854
21.6942 | -105.7556
21.6979 | -105.8254
21.7017 | -105.8931
21.7054 | -105.9585
21.7092 | -106.0232
21.7129 | -106.0846
21.7167 | -106.142
21.7205 | -106.1978
21.7242 | -106.2474
21.728 | -106.2911
21.7317 | -106.3318
21.7355 | -106.365
21.7393 | -106.3946
21.743 | -106.4155
21.7468 | -106.4355
21.7505 | -106.4601
21.7543 | -106.4771
21.758 | -106.4936
21.7618 | -106.5075
21.7656 | -106.5126
21.7693 | -106.5195
21.7731 | -106.5226
21.7768 | -106.5198
21.7806 | -106.5178
21.7844 | -106.516
21.7881 | -106.5098
21.7919 | -106.5005
21.7956 | -106.4923
21.7994 | -106.4857
21.8031 | -106.4748
21.8069 | -106.4623
21.8107 | -106.4508
21.8144 | -106.4395
21.8182 | -106.4271
21.8219 | -106.4121
21.8257 | -106.3984
21.8295 | -106.3867
21.8332 | -106.376
21.837 | -106.3635
21.8407 | -106.3506
21.8445 | -106.3375
21.8483 | -106.3261
21.852 | -106.316
21.8558 | -106.305
21.8595 | -106.2946
21.8633 | -106.2846
21.867 | -106.2751
21.8708 | -106.2647
21.8746 | -106.2544
21.8783 | -106.2457
21.8821 | -106.2374
21.8858 | -106.2302
21.8896 | -106.2252
21.8934 | -106.2202
21.8971 | -106.2147
21.9009 | -106.2122
21.9046 | -106.209
21.9084 | -106.2066
21.9121 | -106.207
21.9159 | -106.2068
21.9197 | -106.2061
21.9234 | -106.2074
21.9272 | -106.2115
21.9309 | -106.2151
21.9347 | -106.2196
21.9385 | -106.2254
21.9422 | -106.2307
21.946 | -106.234
21.9497 | -106.2371
21.9535 | -106.2403
21.9572 | -106.2446
21.961 | -106.2503
21.9648 | -106.2567
21.9685 | -106.263
21.9723 | -106.2678
21.976 | -106.2721
21.9798 | -106.277
21.9836 | -106.2815
21.9873 | -106.2859
21.9911 | -106.2874
21.9948 | -106.2907
21.9986 | -106.2966
22.0023 | -106.3018
22.0061 | -106.309
22.0099 | -106.3146
22.0136 | -106.3186
22.0174 | -106.3242
22.0211 | -106.3287
22.0249 | -106.3309
22.0287 | -106.3356
22.0324 | -106.3394
22.0362 | -106.3419
22.0399 | -106.345
22.0437 | -106.3467
22.0475 | -106.3482
22.0512 | -106.35
22.055 | -106.3539
22.0587 | -106.3554
22.0625 | -106.3537
22.0662 | -106.3534
22.07 | -106.3535
22.0738 | -106.3527
22.0775 | -106.3514
22.0813 | -106.3505
22.085 | -106.3487
22.0888 | -106.3481
22.0926 | -106.3462
22.0963 | -106.3417
22.1001 | -106.3392
22.1038 | -106.3349
22.1076 | -106.3276
22.1113 | -106.3216
22.1151 | -106.3194
22.1189 | -106.3159
22.1226 | -106.3119
22.1264 | -106.3064
22.1301 | -106.3021
22.1339 | -106.2994
22.1377 | -106.295
22.1414 | -106.289
22.1452 | -106.2828
22.1489 | -106.2777
22.1527 | -106.272
22.1564 | -106.2661
22.1602 | -106.2603
22.164 | -106.2552
22.1677 | -106.252
22.1715 | -106.2478
22.1752 | -106.2448
22.179 | -106.2422
22.1828 | -106.2387
22.1865 | -106.2348
22.1903 | -106.2303
22.194 | -106.2249
22.1978 | -106.2166
22.2016 | -106.2096
22.2053 | -106.2044
22.2091 | -106.2012
22.2128 | -106.1958
22.2166 | -106.19
22.2203 | -106.1841
22.2241 | -106.1782
22.2279 | -106.1741
22.2316 | -106.1653
22.2354 | -106.1574
22.2391 | -106.1492
22.2429 | -106.1405
22.2467 | -106.1324
22.2504 | -106.1233
22.2542 | -106.1144
22.2579 | -106.106
22.2617 | -106.0974
22.2654 | -106.0885
22.2692 | -106.0795
22.273 | -106.0712
22.2767 | -106.063
22.2805 | -106.0544
22.2842 | -106.0451
22.288 | -106.0372
22.2918 | -106.0309
22.2955 | -106.0246
22.2993 | -106.019
22.303 | -106.0129
22.3068 | -106.0072
22.3105 | -106.0011
22.3143 | -105.994
22.3181 | -105.9862
22.3218 | -105.9775
22.3256 | -105.97
22.3293 | -105.9625
22.3331 | -105.9553
22.3369 | -105.9513
22.3406 | -105.949
22.3444 | -105.9451
22.3481 | -105.9418
22.3519 | -105.942
22.3557 | -105.9424
22.3594 | -105.9439
22.3632 | -105.9448
22.3669 | -105.9455
22.3707 | -105.9472
22.3744 | -105.95
22.3782 | -105.955
22.382 | -105.9607
22.3857 | -105.9674
22.3895 | -105.975
22.3932 | -105.9846
22.397 | -105.995
22.4008 | -106.0043
22.4045 | -106.0121
22.4083 | -106.0213
22.412 | -106.0311
22.4158 | -106.0425
22.4195 | -106.0534
22.4233 | -106.0621
22.4271 | -106.0691
22.4308 | -106.0752
22.4346 | -106.0826
22.4383 | -106.0874
22.4421 | -106.0823
22.4459 | -106.0683
22.4496 | -106.0498
22.4534 | -106.0308
22.4571 | -106.0105
22.4609 | -105.9917
22.4646 | -105.9748
22.4684 | -105.9581
22.4722 | -105.9396
22.4759 | -105.9192
22.4797 | -105.8938
22.4834 | -105.8652
22.4872 | -105.8308
22.491 | -105.7946
22.4947 | -105.7563
22.4985 | -105.7159
22.5022 | -105.6734
22.506 | -105.6279
22.5097 | -105.5809
22.5135 | -105.5343
22.5173 | -105.4839
22.521 | -105.4306
22.5248 | -105.3753
22.5285 | -105.3201
22.5323 | -105.2655
22.5361 | -105.2118
22.5398 | -105.1571
22.5436 | -105.102
22.5473 | -105.0503
22.5511 | -105.0017
22.5549 | -104.9547
22.5586 | -104.9091
22.5624 | -104.8657
22.5661 | -104.8244
22.5699 | -104.7861
22.5736 | -104.7507
22.5774 | -104.7189
22.5812 | -104.6913
22.5849 | -104.6685
22.5887 | -104.6492
22.5924 | -104.6336
22.5962 | -104.6231
22.6 | -104.6171
22.6037 | -104.6154
22.6075 | -104.6184
22.6112 | -104.6262
22.615 | -104.6388
22.6187 | -104.6551
22.6225 | -104.6767
22.6263 | -104.7036
22.63 | -104.7353
22.6338 | -104.7709
22.6375 | -104.8104
22.6413 | -104.8533
22.6451 | -104.9007
22.6488 | -104.9515
22.6526 | -105.0051
22.6563 | -105.0607
22.6601 | -105.1186
22.6638 | -105.1783
22.6676 | -105.2409
22.6714 | -105.3039
22.6751 | -105.3691
22.6789 | -105.4348
22.6826 | -105.5007
22.6864 | -105.5677
22.6902 | -105.6353
22.6939 | -105.7039
22.6977 | -105.7715
22.7014 | -105.8388
22.7052 | -105.9049
22.709 | -105.9711
22.7127 | -106.0352
22.7165 | -106.096
22.7202 | -106.1564
22.724 | -106.2178
22.7277 | -106.2761
22.7315 | -106.3321
22.7353 | -106.3843
22.739 | -106.4356
22.7428 | -106.4862
22.7465 | -106.5323
22.7503 | -106.5751
22.7541 | -106.6145
22.7578 | -106.655
22.7616 | -106.7012
22.7653 | -106.7484
22.7691 | -106.7932
22.7728 | -106.8027
22.7766 | -106.8384
22.7804 | -106.8701
22.7841 | -106.8936
22.7879 | -106.9081
22.7916 | -106.9181
22.7954 | -106.9274
22.7992 | -106.9345
22.8029 | -106.9438
22.8067 | -106.9531
22.8104 | -106.9639
22.8142 | -106.9712
22.8179 | -106.9775
22.8217 | -106.9838
22.8255 | -106.9898
22.8292 | -106.9953
22.833 | -106.9989
22.8367 | -107.0014
22.8405 | -107.0034
22.8443 | -107.0048
22.848 | -107.0073
22.8518 | -107.0097
22.8555 | -107.0109
22.8593 | -107.0119
22.863 | -107.0127
22.8668 | -107.0139
22.8706 | -107.015
22.8743 | -107.0145
22.8781 | -107.0137
22.8818 | -107.0157
22.8856 | -107.0179
22.8894 | -107.0218
22.8931 | -107.0264
22.8969 | -107.0295
22.9006 | -107.0319
22.9044 | -107.0344
22.9082 | -107.0388
22.9119 | -107.0439
22.9157 | -107.0469
22.9194 | -107.0483
22.9232 | -107.05
22.9269 | -107.0556
22.9307 | -107.061
22.9345 | -107.0652
22.9382 | -107.0697
22.942 | -107.0763
22.9457 | -107.082
22.9495 | -107.0879
22.9533 | -107.092
22.957 | -107.0976
22.9608 | -107.1039
22.9645 | -107.1108
22.9683 | -107.1172
22.972 | -107.1228
22.9758 | -107.129
22.9796 | -107.1343
22.9833 | -107.1395
22.9871 | -107.1456
22.9908 | -107.1496
22.9946 | -107.1543
22.9984 | -107.1583
23.0021 | -107.1612
23.0059 | -107.1654
23.0096 | -107.1684
23.0134 | -107.1686
23.0171 | -107.1688
23.0209 | -107.1673
23.0247 | -107.1655
23.0284 | -107.1653
23.0322 | -107.1627
23.0359 | -107.1603
23.0397 | -107.1592
23.0435 | -107.1581
23.0472 | -107.156
23.051 | -107.1532
23.0547 | -107.1515
23.0585 | -107.1492
23.0622 | -107.146
23.066 | -107.1433
23.0698 | -107.1413
23.0735 | -107.1385
23.0773 | -107.1342
23.081 | -107.129
23.0848 | -107.1252
23.0886 | -107.1222
23.0923 | -107.1177
23.0961 | -107.1124
23.0998 | -107.1091
23.1036 | -107.1062
23.1074 | -107.1
23.1111 | -107.0929
23.1149 | -107.0892
23.1186 | -107.085
23.1224 | -107.0793
23.1261 | -107.076
23.1299 | -107.0706
23.1337 | -107.0652
23.1374 | -107.0601
23.1412 | -107.0577
23.1449 | -107.056
23.1487 | -107.0517
23.1525 | -107.0476
23.1562 | -107.0438
23.16 | -107.0391
23.1637 | -107.0354
23.1675 | -107.0313
23.1712 | -107.0302
23.175 | -107.0271
23.1788 | -107.0224
23.1825 | -107.0192
23.1863 | -107.0159
23.19 | -107.0121
23.1938 | -107.0083
23.1976 | -107.0046
23.2013 | -107.0033
23.2051 | -107.0004
23.2088 | -106.9959
23.2126 | -106.9941
23.2163 | -106.9912
23.2201 | -106.9865
23.2239 | -106.9844
23.2276 | -106.9824
23.2314 | -106.9811
23.2351 | -106.9774
23.2389 | -106.9732
23.2427 | -106.9719
23.2464 | -106.971
23.2502 | -106.9706
23.2539 | -106.9676
23.2577 | -106.9634
23.2615 | -106.9604
23.2652 | -106.9586
23.269 | -106.9565
23.2727 | -106.9544
23.2765 | -106.9512
23.2802 | -106.9511
23.284 | -106.9496
23.2878 | -106.9466
23.2915 | -106.9451
23.2953 | -106.9431
23.299 | -106.9441
23.3028 | -106.944
23.3066 | -106.9452
23.3103 | -106.9463
23.3141 | -106.9503
23.3178 | -106.9503
23.3216 | -106.951
23.3253 | -106.9522
23.3291 | -106.9523
23.3329 | -106.954
23.3366 | -106.9572
23.3404 | -106.96
23.3441 | -106.9627
23.3479 | -106.9656
23.3517 | -106.9696
23.3554 | -106.9727
23.3592 | -106.9738
23.3629 | -106.9761
23.3667 | -106.9787
23.3704 | -106.9817
23.3742 | -106.986
23.378 | -106.9893
23.3817 | -106.9924
23.3855 | -106.9961
23.3892 | -106.9997
23.393 | -107.0046
23.3968 | -107.0097
23.4005 | -107.0134
23.4043 | -107.0183
23.408 | -107.0231
23.4118 | -107.0284
23.4156 | -107.0338
23.4193 | -107.0395
23.4231 | -107.0463
23.4268 | -107.0526
23.4306 | -107.0592
23.4343 | -107.0655
23.4381 | -107.0736
23.4419 | -107.0829
23.4456 | -107.0901
23.4494 | -107.0972
23.4531 | -107.1037
23.4569 | -107.108
23.4607 | -107.1159
23.4644 | -107.1246
23.4682 | -107.1323
23.4719 | -107.1391
23.4757 | -107.144
23.4794 | -107.1502
23.4832 | -107.158
23.487 | -107.1661
23.4907 | -107.1752
23.4945 | -107.1839
23.4982 | -107.1924
23.502 | -107.1989
23.5058 | -107.2073
23.5095 | -107.2162
23.5133 | -107.2237
23.517 | -107.2316
23.5208 | -107.2396
23.5245 | -107.2452
23.5283 | -107.2505
23.5321 | -107.2557
23.5358 | -107.263
23.5396 | -107.2695
23.5433 | -107.2734
23.5471 | -107.2787
23.5509 | -107.2868
23.5546 | -107.292
23.5584 | -107.2969
23.5621 | -107.3028
23.5659 | -107.3083
23.5696 | -107.3159
23.5734 | -107.321
23.5772 | -107.3245
23.5809 | -107.3271
23.5847 | -107.3309
23.5884 | -107.3366
23.5922 | -107.3405
23.596 | -107.3439
23.5997 | -107.3467
23.6035 | -107.3491
23.6072 | -107.3509
23.611 | -107.3526
23.6148 | -107.3547
23.6185 | -107.3588
23.6223 | -107.3609
23.626 | -107.3617
23.6298 | -107.3616
23.6335 | -107.3627
23.6373 | -107.3634
23.6411 | -107.3649
23.6448 | -107.3646
23.6486 | -107.3647
23.6523 | -107.3671
23.6561 | -107.3689
23.6599 | -107.3699
23.6636 | -107.3717
23.6674 | -107.3727
23.6711 | -107.3726
23.6749 | -107.3719
23.6786 | -107.3734
23.6824 | -107.3717
23.6862 | -107.372
23.6899 | -107.3727
23.6937 | -107.3725
23.6974 | -107.3737
23.7012 | -107.3747
23.705 | -107.375
23.7087 | -107.3758
23.7125 | -107.378
23.7162 | -107.3802
23.72 | -107.3802
23.7237 | -107.3821
23.7275 | -107.3841
23.7313 | -107.385
23.735 | -107.3852
23.7388 | -107.3847
23.7425 | -107.3861
23.7463 | -107.3876
23.7501 | -107.3883
23.7538 | -107.3899
23.7576 | -107.3923
23.7613 | -107.3944
23.7651 | -107.3951
23.7689 | -107.3955
23.7726 | -107.3974
23.7764 | -107.3987
23.7801 | -107.3975
23.7839 | -107.3979
23.7876 | -107.4014
23.7914 | -107.4051
23.7952 | -107.4069
23.7989 | -107.4082
23.8027 | -107.4107
23.8064 | -107.4144
23.8102 | -107.4193
23.814 | -107.426
23.8177 | -107.4306
23.8215 | -107.4355
23.8252 | -107.442
23.829 | -107.4463
23.8327 | -107.45
23.8365 | -107.4555
23.8403 | -107.4608
23.844 | -107.4635
23.8478 | -107.4686
23.8515 | -107.4751
23.8553 | -107.483
23.8591 | -107.4892
23.8628 | -107.4982
23.8666 | -107.5067
23.8703 | -107.5139
23.8741 | -107.5193
23.8778 | -107.5264
23.8816 | -107.5346
23.8854 | -107.5412
23.8891 | -107.5479
23.8929 | -107.5552
23.8966 | -107.5612
23.9004 | -107.5679
23.9042 | -107.5742
23.9079 | -107.5787
23.9117 | -107.5837
23.9154 | -107.59
23.9192 | -107.5952
23.923 | -107.5976
23.9267 | -107.5994
23.9305 | -107.6013
23.9342 | -107.6021
23.938 | -107.6041
23.9417 | -107.6045
23.9455 | -107.6038
23.9493 | -107.6019
23.953 | -107.5998
23.9568 | -107.595
23.9605 | -107.591
23.9643 | -107.588
23.9681 | -107.5826
23.9718 | -107.5758
23.9756 | -107.5692
23.9793 | -107.5634
23.9831 | -107.5582
23.9868 | -107.5524
23.9906 | -107.5455
23.9944 | -107.5388
23.9981 | -107.5313
24.0019 | -107.5232
24.0056 | -107.516
24.0094 | -107.5061
24.0132 | -107.4978
24.0169 | -107.4887
24.0207 | -107.4798
24.0244 | -107.4693
24.0282 | -107.4594
24.0319 | -107.4496
24.0357 | -107.4385
24.0395 | -107.4273
24.0432 | -107.4164
24.047 | -107.4045
24.0507 | -107.3918
24.0545 | -107.3814
24.0583 | -107.3717
24.062 | -107.3601
24.0658 | -107.3496
24.0695 | -107.3385
24.0733 | -107.3304
24.077 | -107.322
24.0808 | -107.3133
24.0846 | -107.3054
24.0883 | -107.2977
24.0921 | -107.2886
24.0958 | -107.2806
24.0996 | -107.2735
24.1034 | -107.2677
24.1071 | -107.2639
24.1109 | -107.2585
24.1146 | -107.2538
24.1184 | -107.2527
24.1222 | -107.2508
24.1259 | -107.2478
24.1297 | -107.2465
24.1334 | -107.245
24.1372 | -107.2436
24.1409 | -107.2437
24.1447 | -107.243
24.1485 | -107.2452
24.1522 | -107.2459
24.156 | -107.2462
24.1597 | -107.2459
24.1635 | -107.2462
24.1673 | -107.2508
24.171 | -107.2555
24.1748 | -107.2587
24.1785 | -107.2617
24.1823 | -107.2673
24.186 | -107.2708
24.1898 | -107.2722
24.1936 | -107.2737
24.1973 | -107.2751
24.2011 | -107.2763
24.2048 | -107.2778
24.2086 | -107.2754
24.2124 | -107.274
24.2161 | -107.2737
24.2199 | -107.2705
24.2236 | -107.2682
24.2274 | -107.2679
24.2311 | -107.2659
24.2349 | -107.2641
24.2387 | -107.2625
24.2424 | -107.2586
24.2462 | -107.2533
24.2499 | -107.2481
24.2537 | -107.2356
24.2575 | -107.2086
24.2612 | -107.1717
24.265 | -107.108
24.2687 | -107.0544
24.2725 | -106.9963
24.2763 | -106.9187
24.28 | -106.8062
24.2838 | -106.687
24.2875 | -106.5916
24.2913 | -106.5131
24.295 | -106.4316
24.2988 | -106.3489
24.3026 | -106.2623
24.3063 | -106.1709
24.3101 | -106.0756
24.3138 | -105.9769
24.3176 | -105.8753
24.3214 | -105.7693
24.3251 | -105.6603
24.3289 | -105.5485
24.3326 | -105.4336
24.3364 | -105.3162
24.3401 | -105.198
24.3439 | -105.0796
24.3477 | -104.9608
24.3514 | -104.8424
24.3552 | -104.7244
24.3589 | -104.6079
24.3627 | -104.4932
24.3665 | -104.3793
24.3702 | -104.2675
24.374 | -104.1595
24.3777 | -104.0552
24.3815 | -103.9543
24.3852 | -103.8571
24.389 | -103.7647
24.3928 | -103.6769
24.3965 | -103.5921
24.4003 | -103.548
24.404 | -103.4778
24.4078 | -103.4134
24.4116 | -103.3552
24.4153 | -103.3035
24.4191 | -103.2575
24.4228 | -103.2179
24.4266 | -103.1842
24.4304 | -103.1567
24.4341 | -103.1364
24.4379 | -103.123
24.4416 | -103.1165
24.4454 | -103.1165
24.4491 | -103.1228
24.4529 | -103.136
24.4567 | -103.1564
24.4604 | -103.1827
24.4642 | -103.2155
24.4679 | -103.255
24.4717 | -103.3006
24.4755 | -103.3523
24.4792 | -103.4106
24.483 | -103.4751
24.4867 | -103.5451
24.4905 | -103.6208
24.4942 | -103.7002
24.498 | -103.7846
24.5018 | -103.8744
24.5055 | -103.9688
24.5093 | -104.0667
24.513 | -104.1675
24.5168 | -104.2719
24.5206 | -104.3794
24.5243 | -104.4894
24.5281 | -104.6015
24.5318 | -104.7153
24.5356 | -104.8304
24.5393 | -104.9462
24.5431 | -105.0625
24.5469 | -105.1781
24.5506 | -105.2923
24.5544 | -105.4054
24.5581 | -105.5148
24.5619 | -105.617
24.5657 | -105.713
24.5694 | -105.7966
24.5732 | -105.8715
24.5769 | -105.9314
24.5807 | -105.9724
24.5844 | -106.0014
24.5882 | -106.0011
24.592 | -105.997
24.5957 | -106.0084
24.5995 | -106.0247
24.6032 | -106.0402
24.607 | -106.0496
24.6108 | -106.0619
24.6145 | -106.0467
24.6183 | -106.0201
24.622 | -105.994
24.6258 | -105.9754
24.6296 | -105.9433
24.6333 | -105.8737
24.6371 | -105.775
24.6408 | -105.67
24.6446 | -105.5625
24.6483 | -105.4526
24.6521 | -105.3404
24.6559 | -105.2267
24.6596 | -105.111
24.6634 | -104.994
24.6671 | -104.878
24.6709 | -104.7619
24.6747 | -104.6476
24.6784 | -104.5349
24.6822 | -104.424
24.6859 | -104.3148
24.6897 | -104.2082
24.6934 | -104.1043
24.6972 | -104.0036
24.701 | -103.9065
24.7047 | -103.8129
24.7085 | -103.724
24.7122 | -103.6407
24.716 | -103.5622
24.7198 | -103.4882
24.7235 | -103.4202
24.7273 | -103.3582
24.731 | -103.3021
24.7348 | -103.252
24.7385 | -103.2079
24.7423 | -103.1713
24.7461 | -103.1406
24.7498 | -103.116
24.7536 | -103.0987
24.7573 | -103.088
24.7611 | -103.0834
24.7649 | -103.0859
24.7686 | -103.0948
24.7724 | -103.11
24.7761 | -103.1321
24.7799 | -103.1601
24.7837 | -103.1944
24.7874 | -103.235
24.7912 | -103.2811
24.7949 | -103.3333
24.7987 | -103.3923
24.8024 | -103.4565
24.8062 | -103.5258
24.81 | -103.601
24.8137 | -103.6804
24.8175 | -103.764
24.8212 | -103.8516
24.825 | -103.9434
24.8288 | -104.0388
24.8325 | -104.1376
24.8363 | -104.2393
24.84 | -104.3436
24.8438 | -104.4506
24.8475 | -104.5595
24.8513 | -104.6691
24.8551 | -104.7812
24.8588 | -104.8945
24.8626 | -105.0075
24.8663 | -105.1204
24.8701 | -105.232
24.8739 | -105.3429
24.8776 | -105.4527
24.8814 | -105.5614
24.8851 | -105.6673
24.8889 | -105.7715
24.8926 | -105.8729
24.8964 | -105.9718
24.9002 | -106.0666
24.9039 | -106.159
24.9077 | -106.2485
24.9114 | -106.339
24.9152 | -106.4287
24.919 | -106.5188
24.9227 | -106.604
24.9265 | -106.6916
24.9302 | -106.7858
24.934 | -106.8735
24.9378 | -106.9732
24.9415 | -106.983
24.9453 | -107.0609
24.949 | -107.0981
24.9528 | -107.1185
24.9565 | -107.137
24.9603 | -107.1546
24.9641 | -107.1705
24.9678 | -107.1865
24.9716 | -107.1999
24.9753 | -107.2086
24.9791 | -107.2178
24.9829 | -107.2245
24.9866 | -107.2312
24.9904 | -107.2394
24.9941 | -107.245
24.9979 | -107.2461
25.0016 | -107.248
25.0054 | -107.2493
25.0092 | -107.2491
25.0129 | -107.2486
25.0167 | -107.2474
25.0204 | -107.2451
25.0242 | -107.2428
25.028 | -107.2388
25.0317 | -107.2362
25.0355 | -107.234
25.0392 | -107.2319
25.043 | -107.2286
25.0467 | -107.226
25.0505 | -107.2243
25.0543 | -107.222
25.058 | -107.2212
25.0618 | -107.2198
25.0655 | -107.2174
25.0693 | -107.2157
25.0731 | -107.2149
25.0768 | -107.2106
25.0806 | -107.2064
25.0843 | -107.2032
25.0881 | -107.2013
25.0918 | -107.2007
25.0956 | -107.1987
25.0994 | -107.1987
25.1031 | -107.197
25.1069 | -107.195
25.1106 | -107.1958
25.1144 | -107.1963
25.1182 | -107.1949
25.1219 | -107.1966
25.1257 | -107.1988
25.1294 | -107.1997
25.1332 | -107.2027
25.137 | -107.2056
25.1407 | -107.2095
25.1445 | -107.2143
25.1482 | -107.2179
25.152 | -107.2204
25.1557 | -107.222
25.1595 | -107.2268
25.1633 | -107.2312
25.167 | -107.2364
25.1708 | -107.2431
25.1745 | -107.2499
25.1783 | -107.2579
25.1821 | -107.2646
25.1858 | -107.271
25.1896 | -107.278
25.1933 | -107.284
25.1971 | -107.2902
25.2008 | -107.2969
25.2046 | -107.3037
25.2084 | -107.3081
25.2121 | -107.3107
25.2159 | -107.3142
25.2196 | -107.3172
25.2234 | -107.3211
25.2272 | -107.3232
25.2309 | -107.3249
25.2347 | -107.3258
25.2384 | -107.3281
25.2422 | -107.3288
25.2459 | -107.3292
25.2497 | -107.3275
25.2535 | -107.327
25.2572 | -107.3276
25.261 | -107.3275
25.2647 | -107.3267
25.2685 | -107.325
25.2723 | -107.3235
25.276 | -107.3213
25.2798 | -107.3185
25.2835 | -107.3137
25.2873 | -107.3101
25.291 | -107.3075
25.2948 | -107.3046
25.2986 | -107.3004
25.3023 | -107.2955
25.3061 | -107.29
25.3098 | -107.2849
25.3136 | -107.2775
25.3174 | -107.2715
25.3211 | -107.2688
25.3249 | -107.2666
25.3286 | -107.2624
25.3324 | -107.2586
25.3362 | -107.2532
25.3399 | -107.2485
25.3437 | -107.2441
25.3474 | -107.2399
25.3512 | -107.2356
25.3549 | -107.2299
25.3587 | -107.2216
25.3625 | -107.2131
25.3662 | -107.2077
25.37 | -107.2039
25.3737 | -107.1999
25.3775 | -107.1956
25.3813 | -107.1918
25.385 | -107.1859
25.3888 | -107.1782
25.3925 | -107.172
25.3963 | -107.1659
25.4 | -107.1589
25.4038 | -107.1505
25.4076 | -107.1428
25.4113 | -107.133
25.4151 | -107.1245
25.4188 | -107.1158
25.4226 | -107.1075
25.4264 | -107.0988
25.4301 | -107.0909
25.4339 | -107.0857
25.4376 | -107.079
25.4414 | -107.0708
25.4451 | -107.061
25.4489 | -107.0525
25.4527 | -107.045
25.4564 | -107.0357
25.4602 | -107.0281
25.4639 | -107.0205
25.4677 | -107.0119
25.4715 | -107.0049
25.4752 | -106.9973
25.479 | -106.9874
25.4827 | -106.9828
25.4865 | -106.9755
25.4903 | -106.9694
25.494 | -106.963
25.4978 | -106.9575
25.5015 | -106.9515
25.5053 | -106.9459
25.509 | -106.9404
25.5128 | -106.9343
25.5166 | -106.9305
25.5203 | -106.9261
25.5241 | -106.9222
25.5278 | -106.9191
25.5316 | -106.9165
25.5354 | -106.914
25.5391 | -106.9108
25.5429 | -106.9079
25.5466 | -106.9065
25.5504 | -106.9051
25.5541 | -106.9049
25.5579 | -106.9057
25.5617 | -106.9054
25.5654 | -106.9061
25.5692 | -106.9078
25.5729 | -106.9099
25.5767 | -106.9126
25.5805 | -106.9158
25.5842 | -106.9186
25.588 | -106.9219
25.5917 | -106.9264
25.5955 | -106.9314
25.5992 | -106.9363
25.603 | -106.9417
25.6068 | -106.9477
25.6105 | -106.9546
25.6143 | -106.9604
25.618 | -106.9656
25.6218 | -106.9736
25.6256 | -106.9824
25.6293 | -106.9902
25.6331 | -106.9984
25.6368 | -107.0078
25.6406 | -107.0187
25.6443 | -107.0284
25.6481 | -107.0371
25.6519 | -107.0452
25.6556 | -107.0535
25.6594 | -107.0627
25.6631 | -107.0727
25.6669 | -107.0853
25.6707 | -107.096
25.6744 | -107.1079
25.6782 | -107.1184
25.6819 | -107.1278
25.6857 | -107.1371
25.6895 | -107.1468
25.6932 | -107.1566
25.697 | -107.1668
25.7007 | -107.1762
25.7045 | -107.1863
25.7082 | -107.1964
25.712 | -107.2065
25.7158 | -107.2156
25.7195 | -107.2234
25.7233 | -107.2317
25.727 | -107.2375
25.7308 | -107.2428
25.7346 | -107.25
25.7383 | -107.2577
25.7421 | -107.2655
25.7458 | -107.2735
25.7496 | -107.2826
25.7533 | -107.288
25.7571 | -107.2941
25.7609 | -107.3011
25.7646 | -107.3067
25.7684 | -107.3126
25.7721 | -107.3216
25.7759 | -107.3271
25.7797 | -107.3316
25.7834 | -107.3366
25.7872 | -107.3404
25.7909 | -107.3435
25.7947 | -107.3456
25.7984 | -107.3497
25.8022 | -107.3549
25.806 | -107.3597
25.8097 | -107.3636
25.8135 | -107.3662
25.8172 | -107.3688
25.821 | -107.3715
25.8248 | -107.3733
25.8285 | -107.3739
25.8323 | -107.3736
25.836 | -107.3741
25.8398 | -107.3751
25.8436 | -107.3758
25.8473 | -107.3749
25.8511 | -107.376
25.8548 | -107.3773
25.8586 | -107.3775
25.8623 | -107.3763
25.8661 | -107.3762
25.8699 | -107.3755
25.8736 | -107.375
25.8774 | -107.376
25.8811 | -107.3744
25.8849 | -107.371
25.8887 | -107.368
25.8924 | -107.365
25.8962 | -107.3616
25.8999 | -107.3585
25.9037 | -107.3554
25.9074 | -107.3536
25.9112 | -107.3499
25.915 | -107.3467
25.9187 | -107.3445
25.9225 | -107.3406
25.9262 | -107.3372
25.93 | -107.3343
25.9338 | -107.3313
25.9375 | -107.3282
25.9413 | -107.3239
25.945 | -107.3192
25.9488 | -107.3158
25.9525 | -107.3127
25.9563 | -107.3074
25.9601 | -107.3033
25.9638 | -107.2995
25.9676 | -107.2963
25.9713 | -107.2937
25.9751 | -107.2919
25.9789 | -107.2914
25.9826 | -107.2895
25.9864 | -107.2845
25.9901 | -107.2796
25.9939 | -107.2768
25.9977 | -107.2747
26.0014 | -107.2711
26.0052 | -107.2675
26.0089 | -107.2657
26.0127 | -107.2653
26.0164 | -107.2651
26.0202 | -107.2659
26.024 | -107.2669
26.0277 | -107.2671
26.0315 | -107.2663
26.0352 | -107.2622
26.039 | -107.26
26.0428 | -107.2585
26.0465 | -107.2571
26.0503 | -107.2565
26.054 | -107.2569
26.0578 | -107.2574
26.0615 | -107.2571
26.0653 | -107.2565
26.0691 | -107.2567
26.0728 | -107.2567
26.0766 | -107.2557
26.0803 | -107.2553
26.0841 | -107.2569
26.0879 | -107.259
26.0916 | -107.261
26.0954 | -107.263
26.0991 | -107.2647
26.1029 | -107.2643
26.1066 | -107.2646
26.1104 | -107.2651
26.1142 | -107.2666
26.1179 | -107.2665
26.1217 | -107.2676
26.1254 | -107.2667
26.1292 | -107.2656
26.133 | -107.2677
26.1367 | -107.2692
26.1405 | -107.2698
26.1442 | -107.2704
26.148 | -107.2722
26.1517 | -107.274
26.1555 | -107.2765
26.1593 | -107.2769
26.163 | -107.2774
26.1668 | -107.2774
26.1705 | -107.2754
26.1743 | -107.2756
26.1781 | -107.2759
26.1818 | -107.277
26.1856 | -107.2798
26.1893 | -107.2825
26.1931 | -107.2859
26.1969 | -107.2886
26.2006 | -107.2922
26.2044 | -107.2968
26.2081 | -107.3008
26.2119 | -107.3028
26.2156 | -107.3048
26.2194 | -107.3062
26.2232 | -107.3094
26.2269 | -107.3131
26.2307 | -107.3165
26.2344 | -107.3179
26.2382 | -107.3197
26.242 | -107.3232
26.2457 | -107.3268
26.2495 | -107.3304
26.2532 | -107.3345
26.257 | -107.3384
26.2607 | -107.3415
26.2645 | -107.3439
26.2683 | -107.3468
26.272 | -107.3501
26.2758 | -107.3526
26.2795 | -107.356
26.2833 | -107.36
26.2871 | -107.3634
26.2908 | -107.367
26.2946 | -107.3708
26.2983 | -107.3739
26.3021 | -107.3752
26.3058 | -107.3775
26.3096 | -107.3799
26.3134 | -107.3845
26.3171 | -107.388
26.3209 | -107.3913
26.3246 | -107.393
26.3284 | -107.393
26.3322 | -107.3944
26.3359 | -107.3965
26.3397 | -107.3976
26.3434 | -107.399
26.3472 | -107.4001
26.351 | -107.4008
26.3547 | -107.4037
26.3585 | -107.4061
26.3622 | -107.408
26.366 | -107.4096
26.3697 | -107.4121
26.3735 | -107.4135
26.3773 | -107.414
26.381 | -107.4143
26.3848 | -107.4136
26.3885 | -107.413
26.3923 | -107.413
26.3961 | -107.4128
26.3998 | -107.4132
26.4036 | -107.4149
26.4073 | -107.4188
26.4111 | -107.4214
26.4148 | -107.4212
26.4186 | -107.4195
26.4224 | -107.4186
26.4261 | -107.4189
26.4299 | -107.4202
26.4336 | -107.4218
26.4374 | -107.4231
26.4412 | -107.4245
26.4449 | -107.4239
26.4487 | -107.4251
26.4524 | -107.4268
26.4562 | -107.4274
26.4599 | -107.428
26.4637 | -107.4272
26.4675 | -107.4254
26.4712 | -107.4229
26.475 | -107.4196
26.4787 | -107.4178
26.4825 | -107.4164
26.4863 | -107.4162
26.49 | -107.4178
26.4938 | -107.4193
26.4975 | -107.421
26.5013 | -107.4216
26.5051 | -107.4203
26.5088 | -107.4185
26.5126 | -107.4176
26.5163 | -107.4167
26.5201 | -107.4164
26.5238 | -107.4142
26.5276 | -107.4131
26.5314 | -107.4139
26.5351 | -107.4138
26.5389 | -107.4139
26.5426 | -107.4166
26.5464 | -107.4163
26.5502 | -107.416
26.5539 | -107.4152
26.5577 | -107.4127
26.5614 | -107.412
26.5652 | -107.4116
26.5689 | -107.4091
26.5727 | -107.4081
26.5765 | -107.4067
26.5802 | -107.4037
26.584 | -107.4028
26.5877 | -107.4028
26.5915 | -107.4024
26.5953 | -107.4005
26.599 | -107.3981
26.6028 | -107.3967
26.6065 | -107.3962
26.6103 | -107.3964
26.614 | -107.3957
26.6178 | -107.3944
26.6216 | -107.3918
26.6253 | -107.3892
26.6291 | -107.3868
26.6328 | -107.3838
26.6366 | -107.3806
26.6404 | -107.376
26.6441 | -107.3726
26.6479 | -107.3685
26.6516 | -107.3638
26.6554 | -107.3597
26.6592 | -107.3549
26.6629 | -107.3505
26.6667 | -107.3474
26.6704 | -107.3444
26.6742 | -107.3414
26.6779 | -107.338
26.6817 | -107.3343
26.6855 | -107.3316
26.6892 | -107.3276
26.693 | -107.3215
26.6967 | -107.315
26.7005 | -107.3134
26.7043 | -107.3115
26.708 | -107.308
26.7118 | -107.3032
26.7155 | -107.2982
26.7193 | -107.2938
26.723 | -107.2904
26.7268 | -107.2869
26.7306 | -107.2821
26.7343 | -107.2776
26.7381 | -107.2721
26.7418 | -107.2677
26.7456 | -107.2635
26.7494 | -107.2588
26.7531 | -107.2521
26.7569 | -107.2451
26.7606 | -107.2406
26.7644 | -107.2362
26.7681 | -107.2307
26.7719 | -107.2262
26.7757 | -107.2213
26.7794 | -107.214
26.7832 | -107.2066
26.7869 | -107.2003
26.7907 | -107.1948
26.7945 | -107.189
26.7982 | -107.1819
26.802 | -107.1742
26.8057 | -107.1653
26.8095 | -107.1577
26.8132 | -107.1496
26.817 | -107.1409
26.8208 | -107.1316
26.8245 | -107.1208
26.8283 | -107.1115
26.832 | -107.1025
26.8358 | -107.0933
26.8396 | -107.0866
26.8433 | -107.0797
26.8471 | -107.0719
26.8508 | -107.0651
26.8546 | -107.0559
26.8584 | -107.0469
26.8621 | -107.0378
26.8659 | -107.0286
26.8696 | -107.0193
26.8734 | -107.0099
26.8771 | -107.0029
26.8809 | -106.9952
26.8847 | -106.986
26.8884 | -106.976
26.8922 | -106.9673
26.8959 | -106.9583
26.8997 | -106.9498
26.9035 | -106.9405
26.9072 | -106.932
26.911 | -106.9244
26.9147 | -106.9153
26.9185 | -106.9062
26.9222 | -106.8977
26.926 | -106.8903
26.9298 | -106.884
26.9335 | -106.878
26.9373 | -106.8707
26.941 | -106.8638
26.9448 | -106.8602
26.9486 | -106.8568
26.9523 | -106.852
26.9561 | -106.8454
26.9598 | -106.8404
26.9636 | -106.8357
26.9673 | -106.8324
26.9711 | -106.8305
26.9749 | -106.8304
26.9786 | -106.8279
26.9824 | -106.8272
26.9861 | -106.8276
26.9899 | -106.8274
26.9937 | -106.8262
26.9974 | -106.8249
27.0012 | -106.8258
27.0049 | -106.8275
27.0087 | -106.8291
27.0125 | -106.8323
27.0162 | -106.8353
27.02 | -106.8395
27.0237 | -106.8417
27.0275 | -106.8449
27.0312 | -106.8494
27.035 | -106.8519
27.0388 | -106.8544
27.0425 | -106.8588
27.0463 | -106.8618
27.05 | -106.8643
27.0538 | -106.8669
27.0576 | -106.8716
27.0613 | -106.876
27.0651 | -106.878
27.0688 | -106.8799
27.0726 | -106.8841
27.0763 | -106.8876
27.0801 | -106.8875
27.0839 | -106.8886
27.0876 | -106.8896
27.0914 | -106.8899
27.0951 | -106.8908
27.0989 | -106.8917
27.1027 | -106.8914
27.1064 | -106.8909
27.1102 | -106.8888
27.1139 | -106.8886
27.1177 | -106.8878
27.1214 | -106.8871
27.1252 | -106.8865
27.129 | -106.8865
27.1327 | -106.8859
27.1365 | -106.8865
27.1402 | -106.8876
27.144 | -106.888
27.1478 | -106.888
27.1515 | -106.8894
27.1553 | -106.8899
27.159 | -106.8887
27.1628 | -106.8884
27.1665 | -106.8873
27.1703 | -106.8867
27.1741 | -106.887
27.1778 | -106.8879
27.1816 | -106.8901
27.1853 | -106.8916
27.1891 | -106.8929
27.1929 | -106.8938
27.1966 | -106.896
27.2004 | -106.8999
27.2041 | -106.9064
27.2079 | -106.9132
27.2117 | -106.918
27.2154 | -106.9231
27.2192 | -106.9281
27.2229 | -106.9324
27.2267 | -106.9384
27.2304 | -106.9447
27.2342 | -106.9506
27.238 | -106.9569
27.2417 | -106.9637
27.2455 | -106.9702
27.2492 | -106.9762
27.253 | -106.9838
27.2568 | -106.9905
27.2605 | -106.9965
27.2643 | -107.0047
27.268 | -107.0134
27.2718 | -107.0214
27.2755 | -107.0284
27.2793 | -107.0351
27.2831 | -107.0433
27.2868 | -107.0509
27.2906 | -107.0571
27.2943 | -107.0645
27.2981 | -107.073
27.3019 | -107.0801
27.3056 | -107.088
27.3094 | -107.0959
27.3131 | -107.1023
27.3169 | -107.1078
27.3206 | -107.1145
27.3244 | -107.1214
27.3282 | -107.1272
27.3319 | -107.1329
27.3357 | -107.1403
27.3394 | -107.1492
27.3432 | -107.1568
27.347 | -107.1632
27.3507 | -107.1671
27.3545 | -107.1718
27.3582 | -107.1765
27.362 | -107.1831
27.3658 | -107.1884
27.3695 | -107.1933
27.3733 | -107.1998
27.377 | -107.2077
27.3808 | -107.2139
27.3845 | -107.2183
27.3883 | -107.2233
27.3921 | -107.2292
27.3958 | -107.2348
27.3996 | -107.2408
27.4033 | -107.2456
27.4071 | -107.2514
27.4109 | -107.2564
27.4146 | -107.2617
27.4184 | -107.2671
27.4221 | -107.2721
27.4259 | -107.2759
27.4296 | -107.2815
27.4334 | -107.2859
27.4372 | -107.2898
27.4409 | -107.2947
27.4447 | -107.2993
27.4484 | -107.3054
27.4522 | -107.3113
27.456 | -107.3156
27.4597 | -107.3205
27.4635 | -107.3252
27.4672 | -107.3288
27.471 | -107.3323
27.4747 | -107.3381
27.4785 | -107.3431
27.4823 | -107.3457
27.486 | -107.3469
27.4898 | -107.3477
27.4935 | -107.3493
27.4973 | -107.3518
27.5011 | -107.3543
27.5048 | -107.3572
27.5086 | -107.3599
27.5123 | -107.3605
27.5161 | -107.3604
27.5199 | -107.3609
27.5236 | -107.3608
27.5274 | -107.3585
27.5311 | -107.3574
27.5349 | -107.3552
27.5386 | -107.3516
27.5424 | -107.3486
27.5462 | -107.3478
27.5499 | -107.3455
27.5537 | -107.3445
27.5574 | -107.3428
27.5612 | -107.3404
27.565 | -107.3374
27.5687 | -107.3317
27.5725 | -107.3262
27.5762 | -107.3218
27.58 | -107.3165
27.5837 | -107.3101
27.5875 | -107.3036
27.5913 | -107.2967
27.595 | -107.2896
27.5988 | -107.2821
27.6025 | -107.2735
27.6063 | -107.2657
27.6101 | -107.2571
27.6138 | -107.2474
27.6176 | -107.2393
27.6213 | -107.2318
27.6251 | -107.2234
27.6288 | -107.2134
27.6326 | -107.2037
27.6364 | -107.1957
27.6401 | -107.1869
27.6439 | -107.1782
27.6476 | -107.1674
27.6514 | -107.1539
27.6552 | -107.143
27.6589 | -107.1325
27.6627 | -107.1224
27.6664 | -107.1124
27.6702 | -107.1023
27.6739 | -107.0908
27.6777 | -107.0783
27.6815 | -107.065
27.6852 | -107.0499
27.689 | -107.0351
27.6927 | -107.0201
27.6965 | -107.0044
27.7003 | -106.9906
27.704 | -106.976
27.7078 | -106.9618
27.7115 | -106.9457
27.7153 | -106.9298
27.7191 | -106.9137
27.7228 | -106.898
27.7266 | -106.8815
27.7303 | -106.8667
27.7341 | -106.8524
27.7378 | -106.8362
27.7416 | -106.8202
27.7454 | -106.8047
27.7491 | -106.7892
27.7529 | -106.7724
27.7566 | -106.7562
27.7604 | -106.7383
27.7642 | -106.7217
27.7679 | -106.7043
27.7717 | -106.6884
27.7754 | -106.6734
27.7792 | -106.6589
27.7829 | -106.6435
27.7867 | -106.6274
27.7905 | -106.61
27.7942 | -106.5941
27.798 | -106.5794
27.8017 | -106.564
27.8055 | -106.5478
27.8093 | -106.5331
27.813 | -106.5181
27.8168 | -106.504
27.8205 | -106.4906
27.8243 | -106.4792
27.828 | -106.4678
27.8318 | -106.4546
27.8356 | -106.4425
27.8393 | -106.4309
27.8431 | -106.4192
27.8468 | -106.4083
27.8506 | -106.398
27.8544 | -106.387
27.8581 | -106.3743
27.8619 | -106.3606
27.8656 | -106.3494
27.8694 | -106.3394
27.8731 | -106.3314
27.8769 | -106.3257
27.8807 | -106.3198
27.8844 | -106.3119
27.8882 | -106.304
27.8919 | -106.2965
27.8957 | -106.2886
27.8995 | -106.2823
27.9032 | -106.2769
27.907 | -106.2723
27.9107 | -106.2672
27.9145 | -106.2629
27.9183 | -106.2595
27.922 | -106.2537
27.9258 | -106.241
27.9295 | -106.2197
27.9333 | -106.2045
27.937 | -106.1959
27.9408 | -106.1886
27.9446 | -106.1798
27.9483 | -106.1714
27.9521 | -106.1627
27.9558 | -106.1545
27.9596 | -106.1461
27.9634 | -106.1374
27.9671 | -106.1279
27.9709 | -106.1196
27.9746 | -106.1104
27.9784 | -106.099
27.9821 | -106.0885
27.9859 | -106.0775
27.9897 | -106.0654
27.9934 | -106.0525
27.9972 | -106.0366
28.0009 | -106.0206