## Slide 1: New optical microbarometer
Integrated optics technology
1- Interferometer is designed with
optical waveguide
2 - Optical waveguide are
obtained by an ion-exchange
process into a glass wafer
3 – Interferometer is included
into a compact package
Manufactured by TEEM PHOTONICS
Serge OLIVIER CEA DAM DIF, F-91297
Arpajon, France, serge.olivier@cea.fr
SnT 2015, Vienna Austria
Abstract
Devices under test:
One microbarometer MB2005
Six optical microbarometers:
Optogeo 1, Optogeo 2 and Optogeo 3 used a MB3
aneroid capsule with the corresponding sensitivity
Optogeo 4, Optogeo 5 and Optogeo 6 used a MB2005
aneroid capsule with the corresponding sensitivity
PSD & coherency:
good agreement. Optogeo sensitivity is OK
Noise evaluation:
better than MB2005 for all the bandwidth
Comparison with MB2005
Aneroid capsule area
Optical area
Interferometer
Fiber
optic
Aneroid
capsule
Advantages:
1/ Better resolution for the
total bandwidth
2/ No adjustment with altitude
3/ Less optical adjustment
compared with other
optical technology
Drawbacks:
1/ At present time: price
2/ Fiber optics need lot of space
3/ More losses compared with
other optical technology
First prototype
Digital microbarometer
Analog microbarometer
Digitizer
Mechanical transducer
Movement transducer
Pressure
variation
Displacement
Analog voltage
Digital data
Current microbarometer
Optical microbarometer
Microbarometer principle
Aneroid
capsule
LVDT or
Magnet-coil
Aneroid
capsule
Interferometer
Moving
mirror
Optical
beam
1
3
4
Interferometer principle
2
This project is funded by the DGA though convention n°132906059
Anthony HUE, Nathalie OLIVIER, Serge LE MALLET SEISMO WAVE, Lannion, France, marketing@seismowave.com – www.seismowave.com
CEA DAM (designer of MB series) and PROLANN / SEISMO WAVE (manufacturer and seller of MB3) have associated their expertise to design a new optical microbarometer:
We aim at thinking that changing the electromagnetic transducer by an interferometer is an interesting solution in order to increase the dynamic and the resolution of the sensor.
We propose a future optical microbarometer which will enlarge the panel of infrasound sensors. Results with a first configuration are presented here.
Effects of temperature
5
Axial deformation: measurement principle
Transversal deformation: due to symmetrical errors during manufacturing
Risk: misalignment due to transversal deformation
Measurement: Observation of Optical beam intensity variation with altitude
Effects of altitude
Drawbacks of the first prototype:
Sensitivity of optical measurement to environmental parameters such as humidity
Solution:
To insert interferometer into the aneroid capsule under vacuum
Future design
6
7
Moving mirror
Aneroid
capsule
Alt.= 0 m
Aneroid capsule
Alt.= 2000 m
Optical beam
Moving mirror
Aneroid
capsule
Vacuum
Optical beam
Interferometer
Very low effects at an altitude of 2000 m for sensor with MB3 aneroid capsule
Low effects at an altitude of 2000 m for sensor with MB2005 aneroid sensor
Moving mirror
Interferometer
Aneroid
capsule
Area under
vacuum
Fiber feedthrough
Low thermal effects on infrasound sensitivity compared to MB2005*
* MB2005 thermal drift is less than ±0.1 hPa/°C
Thermal test:
From 5 °C to 40 °C
Step : 5 °C
T3.1-P19