## CALIBRATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 67 |  | CTBTO
Transducer Block | 02-90C/P | 64 |  | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
Hood | 02-92B/P | 53
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE REFERENCE
Atmospheric pressure | 1009.6 | hPa |  | PR 000 150_-_CAL
Humidity | 17.1 | %/Hr |  | Measurement device
Internal Temperature of Calibration Bench | 23.38 | °C |  | IS CALIBRATOR
AUX Sensor pressure | 1010 | hPa |  | Sensor SoH at test start
AUX Sensor Temperature | 23.7 | °C |  | OK
SENSITIVITY AT 1 Hz
Output Parameters | Theoretical Value | Measured Value @1Hz | Margin (±) | Unit | Result
Sensitivity of Pressure Derivative output | 2 | 2 | 0.04 | mV/Pa/sec | OK | Mode d'emploi pour créer un constat de vérification
Sensitivity of Pressure output | 20 | 20 | 0.4 | mV/Pa | OK | Cliquer sur "DEMARRER LA SAISIE"
Sensitivity of Calibration coil | 6 | 5.98 | 0.13 | Pa/V | OK | Suivre les instructions affichées
Une fois la partie SENSOR IDENTIFICATION remplie, renseigner la partie ENVIRONNEMENTAL PARAMETERS (les données environnementales)
TRANSFER FUNCTION | [5 empty] | Renseigner la partie SENSITIVITY AT 1Hz :les sensibilités en voie de dérivée de Pression (Pressure Derivative), en voie de pression (Pressure) et de Calibration (calibration Coil)
Electrical |  | Acoustical |  |  |  | Une sauvegarde temporaire est toujours possible si le fichier doit être modifié (Fichier "enregistrer")
Vérifier toutes les données saisie notamment la partie CUSTOMER ("Client") et la partie VISA (Noms des personnes et date d'ajustage)
Cliquer sur "ENREGISTREMENT DEFINITIF" (Attention le fichier sera enregistré et verrouillé dans le dossier courant, il faut retirer manuellement la protection)
SENSOR SELF NOISE |  |  |  | Decision : considering the specifications given by the manufacturer, this sensor is
CONFORM
VISA
|  |  |  | Adjusted by | A.HUE
|  |  |  | on | 2016-03-17
|  |  |  | Approved by | A.HUE
|  |  |  | Qualification | Optronic Engineer
Self Noise at 1 HZ | 5.6754e-05 | Pa/√Hz |  | on | 2016-03-17 | -84.92 | (dBPa/√Hz)
| < 0,13 | mPa/√Hz
OK
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which define the verification report. It can be used to demonstrate the link between the measure equipment and the standard national or international equipment, provided it meets the recommendations of the documentation booklet X 07-015
[8 empty] | Client | Adresse Client
[8 empty] | CEA | C.E.A. Dam. Bruyères Domaine du Grand Rué 91297 ARPAJON CEDEX FRANCE
[8 empty] | CTBTO | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
[8 empty] | IPGP | Université Paris Diderot Paris 7 Bat Lamarck UFR STEP 35 rue H.Brion 6ème étage 75205 Paris CEDEX 13
[8 empty] | SANDIA | C/O SANDIATIONAL LABS ALBUQUERQUE, NM 87123 UNITED STATES
[8 empty] | ISLA | Haiwaï Institute of Geophysics & Planetology 73-984 Makako Bay Drive KAILUA-KONA HAWAÏ
ENREGISTREMENT DEFINITIF
DEMARRER LA SAISIE

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.1266 | 11.9704
0.097656 | -34.0095 | 7.7641
0.12207 | -33.9409 | 5.5089
0.14648 | -33.9606 | 4.0821
0.1709 | -33.9623 | 3.4293
0.19531 | -33.922 | 2.7482
0.21973 | -33.9262 | 2.5498
0.24414 | -33.9692 | 2.0436
0.26855 | -33.9806 | 1.9814
0.29297 | -33.9396 | 1.1351
0.31738 | -33.9334 | 1.0495
0.3418 | -33.9452 | 0.64005
0.36621 | -33.9908 | 0.8355
0.39062 | -33.9166 | 0.75483
0.41504 | -33.9471 | 0.56759
0.43945 | -33.998 | 0.36969
0.46387 | -33.9254 | -0.079798
0.48828 | -33.9402 | 0.070476
0.5127 | -33.9507 | 0.22424
0.53711 | -33.9938 | -0.16779
0.56152 | -33.9691 | 0.081393
0.58594 | -33.9571 | -0.23451
0.61035 | -33.9827 | -0.33865
0.63477 | -33.9215 | -0.32946
0.65918 | -33.949 | -0.97174
0.68359 | -33.9454 | -0.27441
0.70801 | -33.9676 | -0.69542
0.73242 | -33.9643 | -0.7623
0.75684 | -33.9615 | -0.93322
0.78125 | -33.9816 | -0.37381
0.80566 | -34.0098 | -1.1496
0.83008 | -33.9289 | -0.93085
0.85449 | -33.9194 | -0.72124
0.87891 | -33.9652 | -0.99253
0.90332 | -33.9545 | -0.9993
0.92773 | -34.0037 | -0.6269
0.95215 | -33.9797 | -0.89748
0.97656 | -34.0088 | -1.1924
1.001 | -33.9791 | -1.1298
1.0254 | -33.97 | -1.4066
1.0498 | -34.0116 | -1.6969
1.0742 | -33.974 | -1.561
1.0986 | -33.9694 | -1.6901
1.123 | -33.9715 | -2.0434
1.1475 | -33.9651 | -1.5103
1.1719 | -33.9939 | -1.6937
1.1963 | -34.0118 | -1.74
1.2207 | -33.9608 | -2.1314
1.2451 | -33.9986 | -1.801
1.2695 | -33.9596 | -1.9769
1.2939 | -33.9952 | -2.1221
1.3184 | -34.0424 | -1.947
1.3428 | -34.0205 | -1.558
1.3672 | -34.0323 | -2.3641
1.3916 | -33.9916 | -2.2995
1.416 | -34.0078 | -1.9636
1.4404 | -33.9953 | -2.0876
1.4648 | -33.9458 | -2.0376
1.4893 | -34.0159 | -2.4522
1.5137 | -33.9843 | -2.1209
1.5381 | -33.9884 | -2.4967
1.5625 | -33.9749 | -2.1896
1.5869 | -33.9419 | -2.6142
1.6113 | -34.0153 | -2.6908
1.6357 | -33.9615 | -3.0596
1.6602 | -33.9957 | -2.7575
1.6846 | -33.9742 | -2.5882
1.709 | -34.0238 | -2.6548
1.7334 | -34.0059 | -3.0598
1.7578 | -33.9954 | -2.7791
1.7822 | -33.9701 | -3.2736
1.8066 | -33.9937 | -2.9211
1.8311 | -33.9406 | -2.933
1.8555 | -34.0246 | -3.0501
1.8799 | -33.9573 | -3.0746
1.9043 | -33.9949 | -3.3064
1.9287 | -33.9942 | -3.3198
1.9531 | -34.052 | -3.4527
1.9775 | -33.987 | -2.8495
2.002 | -33.9893 | -3.3873
2.0264 | -33.9992 | -3.2127
2.0508 | -34.0009 | -3.3622
2.0752 | -34.0095 | -3.1014
2.0996 | -34.0158 | -3.3951
2.124 | -34.0167 | -3.4353
2.1484 | -34.0288 | -3.402
2.1729 | -33.9571 | -3.4126
2.1973 | -34.0049 | -3.8598
2.2217 | -34.0271 | -3.6679
2.2461 | -34.0424 | -3.7177
2.2705 | -34.0245 | -3.7049
2.2949 | -33.9677 | -3.8346
2.3193 | -34.0349 | -3.8454
2.3438 | -33.986 | -4.1299
2.3682 | -34.013 | -3.9484
2.3926 | -33.9871 | -3.7421
2.417 | -34.0056 | -4.1244
2.4414 | -34.017 | -3.6266
2.4658 | -34.0402 | -4.4098
2.4902 | -34.0203 | -4.024
2.5146 | -34.0287 | -4.3335
2.5391 | -34.0534 | -4.1379
2.5635 | -34.0259 | -4.3779
2.5879 | -34.101 | -4.5216
2.6123 | -34.0406 | -4.356
2.6367 | -34.0217 | -4.4689
2.6611 | -34.0351 | -4.3545
2.6855 | -34.068 | -4.2574
2.71 | -33.9953 | -4.7136
2.7344 | -34.0323 | -4.5534
2.7588 | -33.999 | -4.4645
2.7832 | -34.032 | -4.8026
2.8076 | -33.9993 | -4.5057
2.832 | -34.0323 | -4.795
2.8564 | -34.036 | -4.6027
2.8809 | -34.0481 | -5.4281
2.9053 | -34.034 | -4.5872
2.9297 | -34.0667 | -4.9816
2.9541 | -34.0539 | -4.7341
2.9785 | -33.9912 | -4.9472
3.0029 | -34.0105 | -4.9088
3.0273 | -34.0534 | -4.7786
3.0518 | -34.0597 | -4.9765
3.0762 | -34.0232 | -5.2025
3.1006 | -34.0673 | -4.9588
3.125 | -34.0543 | -5.4861
3.1494 | -34.0537 | -5.0153
3.1738 | -34.0472 | -5.3449
3.1982 | -34.0449 | -5.4773
3.2227 | -34.0004 | -5.2239
3.2471 | -34.0854 | -5.3909
3.2715 | -34.0094 | -5.2756
3.2959 | -34.0704 | -5.4474
3.3203 | -34.0262 | -5.1555
3.3447 | -34.0873 | -5.4397
3.3691 | -34.0489 | -5.5285
3.3936 | -34.0275 | -5.8767
3.418 | -34.0223 | -5.7286
3.4424 | -34.0851 | -5.6888
3.4668 | -34.0734 | -5.7094
3.4912 | -34.0412 | -5.881
3.5156 | -34.1099 | -5.7579
3.54 | -34.0083 | -5.6897
3.5645 | -34.0223 | -5.9901
3.5889 | -34.0566 | -6.2119
3.6133 | -34.074 | -6.1122
3.6377 | -34.0465 | -6.2021
3.6621 | -34.0861 | -6.0716
3.6865 | -34.0662 | -5.7878
3.7109 | -34.109 | -5.9253
3.7354 | -34.0487 | -6.163
3.7598 | -34.0386 | -6.344
3.7842 | -34.0734 | -6.8496
3.8086 | -34.0688 | -6.9497
3.833 | -34.0735 | -6.5211
3.8574 | -34.1118 | -6.4523
3.8818 | -34.0618 | -6.2216
3.9062 | -34.0342 | -6.698
3.9307 | -34.0852 | -6.5699
3.9551 | -34.0697 | -6.6584
3.9795 | -34.1344 | -6.7412
4.0039 | -34.1196 | -6.6452
4.0283 | -34.1314 | -6.8197
4.0527 | -34.0256 | -6.9897
4.0771 | -34.1385 | -6.8713
4.1016 | -34.067 | -6.9042
4.126 | -34.1142 | -6.9451
4.1504 | -34.0662 | -6.8961
4.1748 | -34.1085 | -6.9795
4.1992 | -34.0514 | -7.1104
4.2236 | -34.0625 | -7.6276
4.248 | -34.0476 | -7.0314
4.2725 | -34.0713 | -7.1957
4.2969 | -34.098 | -6.9893
4.3213 | -34.0705 | -7.4379
4.3457 | -34.0522 | -7.4114
4.3701 | -34.0752 | -7.455
4.3945 | -34.0607 | -7.02
4.4189 | -34.0831 | -7.3802
4.4434 | -34.1093 | -7.3027
4.4678 | -34.083 | -7.3178
4.4922 | -34.1572 | -7.6746
4.5166 | -34.0496 | -7.6237
4.541 | -34.099 | -7.5434
4.5654 | -34.1075 | -7.5136
4.5898 | -34.1079 | -7.7974
4.6143 | -34.0521 | -7.8535
4.6387 | -34.0888 | -7.7201
4.6631 | -34.0905 | -7.8309
4.6875 | -34.141 | -8.0245
4.7119 | -34.0749 | -7.6659
4.7363 | -34.0977 | -7.9571
4.7607 | -34.1176 | -7.7748
4.7852 | -34.1305 | -8.0206
4.8096 | -34.1315 | -7.5936
4.834 | -34.0865 | -8.3325
4.8584 | -34.1335 | -7.7625
4.8828 | -34.1169 | -8.3651
4.9072 | -34.1254 | -8.3095
4.9316 | -34.1355 | -8.4332
4.9561 | -34.1782 | -8.3763
4.9805 | -34.0974 | -8.3142
5.0049 | -34.15 | -8.8847
5.0293 | -34.1253 | -8.4637
5.0537 | -34.1341 | -8.643
5.0781 | -34.081 | -8.356
5.1025 | -34.1423 | -8.3238
5.127 | -34.1287 | -8.4969
5.1514 | -34.0911 | -8.6979
5.1758 | -34.0822 | -8.6538
5.2002 | -34.1179 | -8.9287
5.2246 | -34.087 | -8.609
5.249 | -34.146 | -9.1464
5.2734 | -34.1328 | -8.5985
5.2979 | -34.1623 | -8.9148
5.3223 | -34.1578 | -8.5184
5.3467 | -34.0697 | -8.8557
5.3711 | -34.1231 | -8.6498
5.3955 | -34.1244 | -9.076
5.4199 | -34.1906 | -9.2907
5.4443 | -34.1256 | -8.9351
5.4688 | -34.1189 | -9.211
5.4932 | -34.0778 | -8.7736
5.5176 | -34.1661 | -9.2191
5.542 | -34.0986 | -9.3576
5.5664 | -34.1491 | -9.4155
5.5908 | -34.1484 | -9.3656
5.6152 | -34.1208 | -9.2537
5.6396 | -34.1083 | -9.4409
5.6641 | -34.1199 | -9.8628
5.6885 | -34.1347 | -9.5387
5.7129 | -34.1047 | -9.6901
5.7373 | -34.1519 | -9.0911
5.7617 | -34.1303 | -9.6844
5.7861 | -34.1543 | -9.4401
5.8105 | -34.1294 | -9.1723
5.835 | -34.179 | -9.3323
5.8594 | -34.1341 | -9.8976
5.8838 | -34.1412 | -10.1479
5.9082 | -34.1513 | -9.7688
5.9326 | -34.1463 | -10.0138
5.957 | -34.1625 | -9.6897
5.9814 | -34.2195 | -10.4169
6.0059 | -34.1409 | -9.7781
6.0303 | -34.1182 | -10.347
6.0547 | -34.1197 | -9.7195
6.0791 | -34.1168 | -10.5302
6.1035 | -34.1771 | -10.199
6.1279 | -34.1576 | -10.3587
6.1523 | -34.1903 | -10.0861
6.1768 | -34.1437 | -10.4611
6.2012 | -34.1831 | -10.3156
6.2256 | -34.1684 | -10.442
6.25 | -34.2009 | -10.2721
6.2744 | -34.1659 | -10.541
6.2988 | -34.1805 | -10.5345
6.3232 | -34.1574 | -10.8185
6.3477 | -34.222 | -10.6785
6.3721 | -34.1798 | -10.5492
6.3965 | -34.1476 | -10.5927
6.4209 | -34.185 | -10.558
6.4453 | -34.204 | -10.9379
6.4697 | -34.1713 | -10.6173
6.4941 | -34.2203 | -11.1982
6.5186 | -34.2306 | -10.6277
6.543 | -34.2317 | -10.8166
6.5674 | -34.2079 | -10.6261
6.5918 | -34.1057 | -11.2124
6.6162 | -34.2219 | -11.142
6.6406 | -34.1771 | -10.9398
6.665 | -34.2084 | -10.9089
6.6895 | -34.169 | -11.1471
6.7139 | -34.234 | -11.1111
6.7383 | -34.219 | -11.1295
6.7627 | -34.1971 | -11.2328
6.7871 | -34.1631 | -11.091
6.8115 | -34.2241 | -11.0595
6.8359 | -34.2026 | -11.0807
6.8604 | -34.2099 | -11.4093
6.8848 | -34.2094 | -11.1566
6.9092 | -34.1892 | -11.2017
6.9336 | -34.2363 | -11.5764
6.958 | -34.2057 | -11.6802
6.9824 | -34.2037 | -11.2876
7.0068 | -34.18 | -11.6305
7.0312 | -34.2479 | -11.5138
7.0557 | -34.2064 | -11.6323
7.0801 | -34.2455 | -11.3556
7.1045 | -34.2112 | -11.4626
7.1289 | -34.2641 | -11.6995
7.1533 | -34.2146 | -11.872
7.1777 | -34.2205 | -11.6678
7.2021 | -34.1685 | -11.4351
7.2266 | -34.2461 | -11.8735
7.251 | -34.1934 | -11.917
7.2754 | -34.2627 | -12.4597
7.2998 | -34.2097 | -12.1325
7.3242 | -34.2792 | -11.9874
7.3486 | -34.2571 | -11.9608
7.373 | -34.2238 | -11.8301
7.3975 | -34.2203 | -11.9095
7.4219 | -34.273 | -12.4186
7.4463 | -34.2253 | -11.9255
7.4707 | -34.2241 | -12.5602
7.4951 | -34.2665 | -12.1527
7.5195 | -34.2097 | -12.4318
7.5439 | -34.2747 | -12.0543
7.5684 | -34.2316 | -12.2051
7.5928 | -34.3122 | -12.5205
7.6172 | -34.2273 | -12.1489
7.6416 | -34.2937 | -12.6472
7.666 | -34.2354 | -12.399
7.6904 | -34.3144 | -12.9838
7.7148 | -34.2829 | -12.2767
7.7393 | -34.3389 | -12.8444
7.7637 | -34.2743 | -12.6248
7.7881 | -34.3111 | -13.1385
7.8125 | -34.2197 | -12.64
7.8369 | -34.1867 | -13.0732
7.8613 | -34.2594 | -12.92
7.8857 | -34.2615 | -12.9902
7.9102 | -34.2535 | -12.2675
7.9346 | -34.3009 | -12.7395
7.959 | -34.2649 | -12.6713
7.9834 | -34.2387 | -13.1296
8.0078 | -34.287 | -13.3319
8.0322 | -34.2761 | -13.301
8.0566 | -34.239 | -12.8781
8.0811 | -34.2295 | -13.1351
8.1055 | -34.2765 | -13.6824
8.1299 | -34.2283 | -13.237
8.1543 | -34.2479 | -12.9983
8.1787 | -34.2714 | -13.1699
8.2031 | -34.2811 | -13.2627
8.2275 | -34.2377 | -12.9189
8.252 | -34.3297 | -13.686
8.2764 | -34.2676 | -13.5147
8.3008 | -34.2989 | -13.6327
8.3252 | -34.3269 | -13.4975
8.3496 | -34.3043 | -13.6913
8.374 | -34.2615 | -13.7857
8.3984 | -34.2515 | -13.8528
8.4229 | -34.2881 | -13.7679
8.4473 | -34.239 | -13.9903
8.4717 | -34.3299 | -13.5866
8.4961 | -34.2626 | -13.828
8.5205 | -34.3017 | -13.7265
8.5449 | -34.2552 | -13.6135
8.5693 | -34.3063 | -14.0654
8.5938 | -34.3232 | -13.7348
8.6182 | -34.2867 | -13.9526
8.6426 | -34.2959 | -14.2306
8.667 | -34.2848 | -14.4039
8.6914 | -34.2691 | -14.1549
8.7158 | -34.2798 | -14.0955
8.7402 | -34.2942 | -14.1554
8.7646 | -34.2989 | -14.0944
8.7891 | -34.3117 | -14.3
8.8135 | -34.2752 | -14.3592
8.8379 | -34.3102 | -13.9598
8.8623 | -34.2771 | -14.8294
8.8867 | -34.3647 | -14.4187
8.9111 | -34.3421 | -14.9372
8.9355 | -34.3709 | -14.4716
8.96 | -34.308 | -14.8047
8.9844 | -34.3561 | -14.7255
9.0088 | -34.3641 | -14.7834
9.0332 | -34.3648 | -14.3768
9.0576 | -34.2959 | -14.7125
9.082 | -34.4023 | -15.0583
9.1064 | -34.3729 | -14.6557
9.1309 | -34.364 | -14.86
9.1553 | -34.331 | -15.0595
9.1797 | -34.3348 | -15.2294
9.2041 | -34.3292 | -14.8358
9.2285 | -34.3481 | -15.5406
9.2529 | -34.3574 | -14.7879
9.2773 | -34.3132 | -15.0555
9.3018 | -34.3373 | -15.0279
9.3262 | -34.3693 | -15.3287
9.3506 | -34.3327 | -15.1164
9.375 | -34.3058 | -14.9825
9.3994 | -34.343 | -15.3104
9.4238 | -34.3711 | -15.8749
9.4482 | -34.391 | -15.2521
9.4727 | -34.3468 | -15.7042
9.4971 | -34.4306 | -15.6544
9.5215 | -34.308 | -15.3893
9.5459 | -34.411 | -15.7151
9.5703 | -34.3527 | -15.408
9.5947 | -34.3757 | -15.3626
9.6191 | -34.3359 | -15.4588
9.6436 | -34.3703 | -15.8204
9.668 | -34.3969 | -15.9697
9.6924 | -34.3218 | -15.9599
9.7168 | -34.3856 | -15.6723
9.7412 | -34.3564 | -16.0864
9.7656 | -34.4013 | -15.6767
9.79 | -34.4078 | -15.8698
9.8145 | -34.4707 | -15.6911
9.8389 | -34.3758 | -15.9826
9.8633 | -34.3629 | -15.8194
9.8877 | -34.3596 | -16.0661
9.9121 | -34.4171 | -15.9545
9.9365 | -34.3543 | -16.0504
9.9609 | -34.3828 | -16.0297
9.9854 | -34.4101 | -15.6819
10.0098 | -34.3963 | -16.4294

## ACOUS_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.1733 | 9.8845
0.097656 | -34.1975 | 5.1463
0.12207 | -34.031 | 3.5574
0.14648 | -34.1087 | 1.5906
0.1709 | -34.0124 | 1.3628
0.19531 | -34.0599 | 0.079376
0.21973 | -34.0405 | 0.39408
0.24414 | -34.0669 | -0.84001
0.26855 | -34.0702 | -0.30716
0.29297 | -34.0291 | -1.5505
0.31738 | -34.0621 | -1.0114
0.3418 | -34.0089 | -1.7814
0.36621 | -34.0855 | -1.4167
0.39062 | -33.991 | -2.0762
0.41504 | -34.0985 | -1.684
0.43945 | -33.9999 | -2.1183
0.46387 | -34.0936 | -2.227
0.48828 | -33.9651 | -2.2251
0.5127 | -34.1101 | -2.5467
0.53711 | -34.0023 | -2.3046
0.56152 | -34.0956 | -2.8674
0.58594 | -34.0131 | -2.426
0.61035 | -34.087 | -3.1982
0.63477 | -34.0411 | -2.5205
0.65918 | -34.081 | -3.3679
0.68359 | -34.0445 | -2.6365
0.70801 | -34.0511 | -3.4295
0.73242 | -34.0915 | -2.814
0.75684 | -34.054 | -3.6511
0.78125 | -34.1072 | -2.9676
0.80566 | -34.0325 | -3.7467
0.83008 | -34.1251 | -3.3325
0.85449 | -34.0024 | -3.7828
0.87891 | -34.0925 | -3.5171
0.90332 | -33.9911 | -3.5757
0.92773 | -34.1345 | -3.5687
0.95215 | -34.035 | -3.5958
0.97656 | -34.1325 | -4.0336
1.001 | -34.0251 | -3.7119
1.0254 | -34.1276 | -4.2599
1.0498 | -34.067 | -3.7338
1.0742 | -34.1198 | -4.5874
1.0986 | -34.047 | -3.8974
1.123 | -34.0677 | -4.5639
1.1475 | -34.0725 | -3.8163
1.1719 | -34.0675 | -4.6865
1.1963 | -34.0958 | -3.8388
1.2207 | -34.0527 | -4.8087
1.2451 | -34.1323 | -4.0545
1.2695 | -34.0347 | -4.7143
1.2939 | -34.1367 | -4.2061
1.3184 | -34.009 | -4.5856
1.3428 | -34.1656 | -4.3652
1.3672 | -34.0472 | -4.5743
1.3916 | -34.1464 | -4.7225
1.416 | -34.0104 | -4.6751
1.4404 | -34.1385 | -4.8859
1.4648 | -34.029 | -4.376
1.4893 | -34.1459 | -4.8903
1.5137 | -34.0751 | -4.4058
1.5381 | -34.1448 | -5.2793
1.5625 | -34.1046 | -4.4964
1.5869 | -34.1246 | -5.3698
1.6113 | -34.126 | -4.7393
1.6357 | -34.0773 | -5.6876
1.6602 | -34.1137 | -4.9057
1.6846 | -34.0608 | -5.4885
1.709 | -34.1456 | -4.9687
1.7334 | -34.0296 | -5.5393
1.7578 | -34.173 | -5.2763
1.7822 | -34.0355 | -5.4181
1.8066 | -34.1692 | -5.3268
1.8311 | -34.0397 | -5.3205
1.8555 | -34.2008 | -5.5509
1.8799 | -34.0677 | -5.3127
1.9043 | -34.161 | -5.9023
1.9287 | -34.0422 | -5.3615
1.9531 | -34.1397 | -6.0339
1.9775 | -34.0797 | -5.1491
2.002 | -34.1337 | -6.1006
2.0264 | -34.0978 | -5.315
2.0508 | -34.1052 | -6.0092
2.0752 | -34.1394 | -5.3697
2.0996 | -34.1152 | -6.1956
2.124 | -34.1653 | -5.5096
2.1484 | -34.0919 | -6.4144
2.1729 | -34.1649 | -5.8511
2.1973 | -34.0714 | -6.2007
2.2217 | -34.1735 | -5.984
2.2461 | -34.0718 | -6.1991
2.2705 | -34.2135 | -6.189
2.2949 | -34.083 | -6.2281
2.3193 | -34.1847 | -6.368
2.3438 | -34.0796 | -6.059
2.3682 | -34.1491 | -6.5552
2.3926 | -34.0745 | -5.8983
2.417 | -34.1381 | -6.6658
2.4414 | -34.1104 | -5.8711
2.4658 | -34.1695 | -6.7367
2.4902 | -34.1563 | -5.9938
2.5146 | -34.1374 | -6.9758
2.5391 | -34.1527 | -6.2558
2.5635 | -34.1169 | -6.902
2.5879 | -34.178 | -6.3104
2.6123 | -34.0913 | -6.7917
2.6367 | -34.1806 | -6.434
2.6611 | -34.0885 | -6.8973
2.6855 | -34.2103 | -6.6254
2.71 | -34.071 | -6.8416
2.7344 | -34.2168 | -6.8467
2.7588 | -34.1011 | -6.579
2.7832 | -34.1937 | -7.1443
2.8076 | -34.0852 | -6.6173
2.832 | -34.1976 | -7.1659
2.8564 | -34.1411 | -6.7036
2.8809 | -34.1639 | -7.5298
2.9053 | -34.1172 | -6.7129
2.9297 | -34.1542 | -7.5471
2.9541 | -34.171 | -6.8594
2.9785 | -34.1149 | -7.592
3.0029 | -34.1765 | -6.8585
3.0273 | -34.1165 | -7.5909
3.0518 | -34.2082 | -6.9428
3.0762 | -34.1072 | -7.505
3.1006 | -34.2073 | -7.1475
3.125 | -34.105 | -7.5003
3.1494 | -34.2219 | -7.4152
3.1738 | -34.1138 | -7.5526
3.1982 | -34.2301 | -7.6386
3.2227 | -34.0895 | -7.387
3.2471 | -34.2032 | -7.837
3.2715 | -34.1317 | -7.3418
3.2959 | -34.2189 | -7.7773
3.3203 | -34.1449 | -7.1817
3.3447 | -34.186 | -8.0016
3.3691 | -34.1771 | -7.2914
3.3936 | -34.1821 | -7.9945
3.418 | -34.1896 | -7.4072
3.4424 | -34.1603 | -8.2711
3.4668 | -34.2252 | -7.5294
3.4912 | -34.1469 | -8.3644
3.5156 | -34.2561 | -7.8149
3.54 | -34.1079 | -8.22
3.5645 | -34.2357 | -7.979
3.5889 | -34.1344 | -8.165
3.6133 | -34.2503 | -8.1021
3.6377 | -34.1108 | -8.2087
3.6621 | -34.2329 | -8.3877
3.6865 | -34.1247 | -7.9328
3.7109 | -34.2521 | -8.4127
3.7354 | -34.17 | -8.0191
3.7598 | -34.222 | -8.6605
3.7842 | -34.1721 | -8.1697
3.8086 | -34.1927 | -8.7362
3.833 | -34.1806 | -8.0404
3.8574 | -34.1882 | -8.7025
3.8818 | -34.2348 | -8.087
3.9062 | -34.1734 | -8.8934
3.9307 | -34.219 | -8.0941
3.9551 | -34.1621 | -8.7888
3.9795 | -34.2854 | -8.4535
4.0039 | -34.1508 | -8.7908
4.0283 | -34.2672 | -8.6062
4.0527 | -34.1302 | -8.7767
4.0771 | -34.2769 | -8.8627
4.1016 | -34.1779 | -8.6626
4.126 | -34.2777 | -8.9732
4.1504 | -34.1731 | -8.7207
4.1748 | -34.2645 | -9.2856
4.1992 | -34.187 | -8.744
4.2236 | -34.2233 | -9.4695
4.248 | -34.1823 | -8.6517
4.2725 | -34.1987 | -9.3696
4.2969 | -34.1997 | -8.5734
4.3213 | -34.2206 | -9.4829
4.3457 | -34.2295 | -8.8154
4.3701 | -34.1767 | -9.5476
4.3945 | -34.2504 | -8.9073
4.4189 | -34.1745 | -9.4713
4.4434 | -34.2977 | -9.0122
4.4678 | -34.2017 | -9.4832
4.4922 | -34.2869 | -9.4512
4.5166 | -34.1472 | -9.4029
4.541 | -34.2889 | -9.5019
4.5654 | -34.1706 | -9.2928
4.5898 | -34.283 | -9.7437
4.6143 | -34.1721 | -9.1652
4.6387 | -34.2579 | -9.9169
4.6631 | -34.2094 | -9.2329
4.6875 | -34.2509 | -10.0299
4.7119 | -34.2174 | -9.2599
4.7363 | -34.2299 | -9.7778
4.7607 | -34.2398 | -9.2804
4.7852 | -34.2367 | -10.1422
4.8096 | -34.2736 | -9.4975
4.834 | -34.195 | -10.1776
4.8584 | -34.3091 | -9.4381
4.8828 | -34.1877 | -9.8862
4.9072 | -34.3143 | -9.7142
4.9316 | -34.1916 | -10.1166
4.9561 | -34.3127 | -10.0525
4.9805 | -34.1814 | -9.851
5.0049 | -34.29 | -10.2795
5.0293 | -34.1573 | -9.8879
5.0537 | -34.2775 | -10.4179
5.0781 | -34.1953 | -9.6763
5.1025 | -34.2951 | -10.5251
5.127 | -34.2236 | -9.7928
5.1514 | -34.2883 | -10.6136
5.1758 | -34.256 | -9.8517
5.2002 | -34.2694 | -10.6609
5.2246 | -34.2832 | -9.8463
5.249 | -34.2084 | -10.623
5.2734 | -34.2901 | -9.9301
5.2979 | -34.2322 | -10.7605
5.3223 | -34.3465 | -10.282
5.3467 | -34.2108 | -10.5867
5.3711 | -34.3359 | -10.3929
5.3955 | -34.1915 | -10.6116
5.4199 | -34.3512 | -10.7346
5.4443 | -34.2204 | -10.5552
5.4688 | -34.316 | -11.0173
5.4932 | -34.2 | -10.4596
5.5176 | -34.3046 | -10.9626
5.542 | -34.2296 | -10.5201
5.5664 | -34.3049 | -11.2211
5.5908 | -34.2466 | -10.5577
5.6152 | -34.2844 | -11.3817
5.6396 | -34.2591 | -10.7558
5.6641 | -34.2478 | -11.3893
5.6885 | -34.2933 | -10.617
5.7129 | -34.221 | -11.3846
5.7373 | -34.2995 | -10.7638
5.7617 | -34.2409 | -11.2821
5.7861 | -34.3653 | -10.6403
5.8105 | -34.2151 | -11.2728
5.835 | -34.3847 | -11.1129
5.8594 | -34.2209 | -11.1545
5.8838 | -34.3585 | -11.2171
5.9082 | -34.2422 | -11.204
5.9326 | -34.3555 | -11.6009
5.957 | -34.2502 | -10.9007
5.9814 | -34.3474 | -11.6158
6.0059 | -34.2662 | -11.1327
6.0303 | -34.3217 | -11.7865
6.0547 | -34.2686 | -11.0685
6.0791 | -34.2856 | -11.6553
6.1035 | -34.3011 | -10.8629
6.1279 | -34.2834 | -11.9964
6.1523 | -34.3339 | -11.2715
6.1768 | -34.251 | -12.0025
6.2012 | -34.3362 | -11.4747
6.2256 | -34.2478 | -11.7501
6.25 | -34.3751 | -11.4631
6.2744 | -34.263 | -11.6726
6.2988 | -34.3571 | -11.8063
6.3232 | -34.2574 | -11.5654
6.3477 | -34.3705 | -11.6936
6.3721 | -34.2449 | -11.4734
6.3965 | -34.3648 | -12.0352
6.4209 | -34.2649 | -11.6091
6.4453 | -34.3472 | -12.1495
6.4697 | -34.2992 | -11.4782
6.4941 | -34.3452 | -12.4921
6.5186 | -34.3388 | -11.5403
6.543 | -34.3369 | -12.5531
6.5674 | -34.3659 | -11.6484
6.5918 | -34.3133 | -12.6095
6.6162 | -34.3483 | -11.9376
6.6406 | -34.2548 | -12.4402
6.665 | -34.3981 | -11.985
6.6895 | -34.2712 | -12.2603
6.7139 | -34.3996 | -12.2818
6.7383 | -34.2673 | -12.3643
6.7627 | -34.3752 | -12.4196
6.7871 | -34.2413 | -12.1791
6.8115 | -34.4167 | -12.5153
6.8359 | -34.3097 | -12.4512
6.8604 | -34.393 | -12.7561
6.8848 | -34.2769 | -12.1684
6.9092 | -34.374 | -12.8644
6.9336 | -34.3567 | -12.3162
6.958 | -34.3544 | -13.0655
6.9824 | -34.358 | -12.2423
7.0068 | -34.3477 | -12.9944
7.0312 | -34.3747 | -12.2454
7.0557 | -34.3384 | -12.9453
7.0801 | -34.3975 | -12.2935
7.1045 | -34.3389 | -12.8115
7.1289 | -34.44 | -12.5036
7.1533 | -34.298 | -12.9222
7.1777 | -34.425 | -12.6844
7.2021 | -34.3163 | -12.9468
7.2266 | -34.4572 | -12.9241
7.251 | -34.356 | -12.9493
7.2754 | -34.4234 | -13.0421
7.2998 | -34.2929 | -12.5254
7.3242 | -34.3856 | -13.2658
7.3486 | -34.3442 | -12.6277
7.373 | -34.3917 | -13.3907
7.3975 | -34.3677 | -12.6119
7.4219 | -34.3892 | -13.5832
7.4463 | -34.3722 | -12.8214
7.4707 | -34.3475 | -13.5553
7.4951 | -34.3672 | -12.9394
7.5195 | -34.3445 | -13.468
7.5439 | -34.4191 | -13.1444
7.5684 | -34.3304 | -13.3308
7.5928 | -34.4039 | -13.1839
7.6172 | -34.3056 | -13.3662
7.6416 | -34.4592 | -13.3338
7.666 | -34.3167 | -13.1132
7.6904 | -34.4821 | -13.611
7.7148 | -34.3622 | -13.3101
7.7393 | -34.4551 | -13.794
7.7637 | -34.4015 | -13.2994
7.7881 | -34.4373 | -14.0758
7.8125 | -34.3901 | -13.4209
7.8369 | -34.3918 | -14.1479
7.8613 | -34.3917 | -13.2131
7.8857 | -34.3833 | -14.2782
7.9102 | -34.4427 | -13.3742
7.9346 | -34.3478 | -14.189
7.959 | -34.448 | -13.5321
7.9834 | -34.3576 | -14.1461
8.0078 | -34.4652 | -13.6867
8.0322 | -34.3801 | -14.163
8.0566 | -34.486 | -13.9434
8.0811 | -34.365 | -13.8465
8.1055 | -34.4761 | -13.8361
8.1299 | -34.3313 | -13.8463
8.1543 | -34.4561 | -14.2844
8.1787 | -34.3633 | -13.7523
8.2031 | -34.4219 | -14.1896
8.2275 | -34.3925 | -13.7614
8.252 | -34.4338 | -14.5311
8.2764 | -34.431 | -13.9262
8.3008 | -34.4325 | -14.5558
8.3252 | -34.4447 | -14.0053
8.3496 | -34.427 | -14.6894
8.374 | -34.4785 | -13.8994
8.3984 | -34.4063 | -14.9339
8.4229 | -34.4599 | -14.0476
8.4473 | -34.4108 | -14.6355
8.4717 | -34.4807 | -14.1654
8.4961 | -34.371 | -14.7219
8.5205 | -34.5076 | -14.1318
8.5449 | -34.3669 | -14.6613
8.5693 | -34.5372 | -14.529
8.5938 | -34.3901 | -14.2192
8.6182 | -34.5151 | -14.9321
8.6426 | -34.4105 | -14.2953
8.667 | -34.4791 | -15.0167
8.6914 | -34.4107 | -14.1204
8.7158 | -34.4508 | -15.2309
8.7402 | -34.4277 | -14.3341
8.7646 | -34.4575 | -15.0426
8.7891 | -34.4914 | -14.4571
8.8135 | -34.4379 | -15.3312
8.8379 | -34.5283 | -14.3491
8.8623 | -34.4322 | -14.9397
8.8867 | -34.5413 | -14.808
8.9111 | -34.4066 | -14.7984
8.9355 | -34.5169 | -15.0241
8.96 | -34.4143 | -15.1444
8.9844 | -34.5218 | -15.109
9.0088 | -34.419 | -15.0318
9.0332 | -34.5323 | -15.1551
9.0576 | -34.3824 | -14.9365
9.082 | -34.5493 | -15.3829
9.1064 | -34.4483 | -14.67
9.1309 | -34.5214 | -15.4957
9.1553 | -34.4403 | -14.7929
9.1797 | -34.4871 | -15.6886
9.2041 | -34.4755 | -14.9247
9.2285 | -34.4709 | -15.6865
9.2529 | -34.4617 | -14.7797
9.2773 | -34.4366 | -15.9222
9.3018 | -34.5111 | -14.9233
9.3262 | -34.4703 | -15.5946
9.3506 | -34.5747 | -15.2123
9.375 | -34.4257 | -15.454
9.3994 | -34.5521 | -15.2031
9.4238 | -34.4673 | -15.4628
9.4482 | -34.5962 | -15.2838
9.4727 | -34.4448 | -15.4355
9.4971 | -34.5524 | -15.8484
9.5215 | -34.4372 | -15.4313
9.5459 | -34.5238 | -16.0957
9.5703 | -34.4782 | -15.2318
9.5947 | -34.5147 | -16.1439
9.6191 | -34.4844 | -15.4266
9.6436 | -34.536 | -16.0816
9.668 | -34.5193 | -15.3855
9.6924 | -34.4514 | -16.0671
9.7168 | -34.5252 | -15.4717
9.7412 | -34.4759 | -16.2075
9.7656 | -34.5692 | -15.3446
9.79 | -34.4424 | -16.1644
9.8145 | -34.5675 | -15.5013
9.8389 | -34.4534 | -16.5267
9.8633 | -34.5801 | -15.8109
9.8877 | -34.4419 | -15.9852
9.9121 | -34.5705 | -15.9574
9.9365 | -34.4529 | -16.0937
9.9609 | -34.5985 | -16.1805
9.9854 | -34.4851 | -15.9472
10.0098 | -34.6078 | -16.603

## SEALED_NOISE
dBPa/√Hz | Frequency
0.010072 | -32.7141
0.010846 | -42.5712
0.011621 | -40.0518
0.012396 | -40.8479
0.013171 | -31.5331
0.013945 | -33.2654
0.01472 | -37.3209
0.015495 | -36.9984
0.016269 | -30.1126
0.017044 | -29.3856
0.017819 | -32.1082
0.018594 | -33.4617
0.019368 | -33.3692
0.020143 | -42.6329
0.020918 | -43.691
0.021693 | -47.2419
0.022467 | -44.1857
0.023242 | -40.6979
0.024017 | -39.9361
0.024792 | -39.8698
0.025566 | -39.601
0.026341 | -39.143
0.027116 | -38.9108
0.027891 | -39.722
0.028665 | -40.6183
0.02944 | -42.316
0.030215 | -42.0387
0.030989 | -40.1106
0.031764 | -38.0932
0.032539 | -39.4085
0.033314 | -43.3769
0.034088 | -47.3769
0.034863 | -49.1983
0.035638 | -47.1303
0.036413 | -44.9795
0.037187 | -44.4394
0.037962 | -45.2001
0.038737 | -45.4721
0.039512 | -46.6724
0.040286 | -49.8791
0.041061 | -52.0827
0.041836 | -47.6602
0.042611 | -44.7756
0.043385 | -44.9057
0.04416 | -46.4381
0.044935 | -48.383
0.04571 | -50.6017
0.046484 | -51.1659
0.047259 | -50.2102
0.048034 | -49.5974
0.048808 | -49.3059
0.049583 | -49.308
0.050358 | -50.3385
0.051133 | -52.3131
0.051907 | -54.0207
0.052682 | -54.8511
0.053457 | -55.0335
0.054232 | -54.9678
0.055006 | -53.4232
0.055781 | -50.7355
0.056556 | -48.3592
0.057331 | -47.2139
0.058105 | -47.5355
0.05888 | -48.9874
0.059655 | -50.595
0.06043 | -51.3014
0.061204 | -51.1863
0.061979 | -51.3193
0.062754 | -52.252
0.063528 | -53.6052
0.064303 | -54.1872
0.065078 | -54.335
0.065853 | -55.041
0.066627 | -55.6103
0.067402 | -55.876
0.068177 | -55.978
0.068952 | -55.9637
0.069726 | -55.5233
0.070501 | -54.3679
0.071276 | -52.8756
0.072051 | -51.8648
0.072825 | -51.7523
0.0736 | -52.4644
0.074375 | -53.5508
0.07515 | -54.3729
0.075924 | -54.8232
0.076699 | -55.1698
0.077474 | -55.5608
0.078248 | -56.0996
0.079023 | -56.5647
0.079798 | -56.6408
0.080573 | -56.4403
0.081347 | -56.4065
0.082122 | -56.8776
0.082897 | -57.8764
0.083672 | -59.184
0.084446 | -60.5366
0.085221 | -61.687
0.085996 | -62.2102
0.086771 | -62.1618
0.087545 | -61.717
0.08832 | -60.9271
0.089095 | -59.3308
0.08987 | -58.1557
0.090644 | -57.1988
0.091419 | -56.6541
0.092194 | -56.5624
0.092968 | -56.7352
0.093743 | -56.8719
0.094518 | -56.6004
0.095293 | -55.9515
0.096067 | -55.2977
0.096842 | -54.9468
0.097617 | -55.0476
0.098392 | -55.5758
0.099166 | -56.4317
0.099941 | -57.4851
0.10072 | -58.6704
0.10149 | -59.8657
0.10226 | -60.5412
0.10304 | -60.3897
0.10382 | -59.6853
0.10459 | -58.8774
0.10536 | -58.1955
0.10614 | -57.8516
0.10691 | -57.9386
0.10769 | -58.4279
0.10846 | -59.1237
0.10924 | -59.5093
0.11001 | -59.1232
0.11079 | -58.1906
0.11156 | -57.5558
0.11234 | -57.2705
0.11311 | -57.326
0.11389 | -57.735
0.11466 | -58.4511
0.11544 | -59.3336
0.11621 | -60.2424
0.11699 | -61.0701
0.11776 | -61.7399
0.11854 | -62.1646
0.11931 | -62.3932
0.12008 | -62.587
0.12086 | -62.8428
0.12163 | -63.3262
0.12241 | -63.9959
0.12318 | -64.778
0.12396 | -65.3148
0.12473 | -65.3396
0.12551 | -64.9104
0.12628 | -64.2353
0.12706 | -63.491
0.12783 | -62.925
0.12861 | -62.5536
0.12938 | -62.2969
0.13016 | -62.1162
0.13093 | -61.9881
0.1317 | -61.9513
0.13248 | -62.0675
0.13326 | -62.3729
0.13403 | -62.957
0.1348 | -63.4917
0.13558 | -64.057
0.13635 | -64.517
0.13713 | -64.8579
0.1379 | -65.1221
0.13868 | -65.207
0.13945 | -65.2322
0.14023 | -65.393
0.141 | -65.5805
0.14178 | -65.7736
0.14255 | -65.8878
0.14333 | -65.896
0.1441 | -65.8745
0.14488 | -65.8483
0.14565 | -65.8776
0.14642 | -65.9764
0.1472 | -66.1218
0.14797 | -66.331
0.14875 | -66.5934
0.14952 | -66.923
0.1503 | -67.3246
0.15107 | -67.8378
0.15185 | -68.2369
0.15262 | -68.5201
0.1534 | -68.5071
0.15417 | -68.2882
0.15495 | -67.9909
0.15572 | -67.9538
0.1565 | -67.9538
0.15727 | -68.1017
0.15805 | -68.3644
0.15882 | -68.73
0.1596 | -69.1327
0.16037 | -69.5178
0.16115 | -69.8373
0.16192 | -70.0677
0.1627 | -70.2423
0.16347 | -70.3245
0.16424 | -70.3869
0.16502 | -70.5473
0.16579 | -70.8409
0.16657 | -71.1969
0.16734 | -71.3595
0.16812 | -71.3722
0.16889 | -71.1778
0.16967 | -70.7122
0.17044 | -70.1243
0.17122 | -69.5337
0.17199 | -69.0396
0.17277 | -68.7142
0.17354 | -68.6017
0.17432 | -68.6973
0.17509 | -68.9684
0.17586 | -69.3678
0.17664 | -69.769
0.17741 | -69.9608
0.17819 | -69.5722
0.17896 | -69.3705
0.17974 | -69.0234
0.18051 | -68.6293
0.18129 | -68.3275
0.18206 | -68.072
0.18284 | -67.9488
0.18361 | -68.0023
0.18439 | -68.2328
0.18516 | -68.6271
0.18594 | -69.1426
0.18671 | -69.7699
0.18749 | -70.4282
0.18826 | -71.0765
0.18904 | -71.6972
0.18981 | -72.1778
0.19059 | -72.5716
0.19136 | -72.9428
0.19214 | -73.0781
0.19291 | -73.0833
0.19368 | -73.0908
0.19446 | -73.1014
0.19523 | -72.9613
0.19601 | -72.5768
0.19678 | -72.1918
0.19756 | -71.6617
0.19833 | -71.0162
0.19911 | -70.3308
0.19988 | -69.6753
0.20066 | -69.1491
0.20143 | -68.7506
0.20221 | -68.4672
0.20298 | -68.3142
0.20376 | -68.2976
0.20453 | -68.4052
0.2053 | -68.6444
0.20608 | -69.0023
0.20686 | -69.4416
0.20763 | -69.9248
0.2084 | -70.4259
0.20918 | -70.8436
0.20995 | -71.1695
0.21073 | -71.4114
0.2115 | -71.5985
0.21228 | -71.6509
0.21305 | -71.7007
0.21383 | -71.7693
0.2146 | -71.8537
0.21538 | -71.9323
0.21615 | -72.0155
0.21693 | -72.0716
0.2177 | -72.1635
0.21848 | -72.2403
0.21925 | -72.2758
0.22002 | -72.3554
0.2208 | -72.4428
0.22157 | -72.5161
0.22235 | -72.5309
0.22312 | -72.5802
0.2239 | -72.6292
0.22467 | -72.7029
0.22545 | -72.8262
0.22622 | -72.9948
0.227 | -73.2274
0.22777 | -73.5317
0.22855 | -73.8603
0.22932 | -74.0959
0.2301 | -74.2808
0.23087 | -74.4192
0.23165 | -74.4045
0.23242 | -74.2044
0.2332 | -73.7886
0.23397 | -73.2745
0.23475 | -72.6576
0.23552 | -72.0322
0.2363 | -71.511
0.23707 | -71.0123
0.23784 | -70.5731
0.23862 | -70.2368
0.23939 | -69.9838
0.24017 | -69.8026
0.24094 | -69.6722
0.24172 | -69.6163
0.24249 | -69.6241
0.24327 | -69.6894
0.24404 | -69.8247
0.24482 | -70.0127
0.24559 | -70.3235
0.24637 | -70.5849
0.24714 | -70.8958
0.24792 | -71.2394
0.24869 | -71.551
0.24946 | -71.7883
0.25024 | -71.9982
0.25101 | -72.1497
0.25179 | -72.233
0.25256 | -72.2197
0.25334 | -72.1147
0.25411 | -71.9976
0.25489 | -71.8364
0.25566 | -71.6887
0.25644 | -71.5824
0.25721 | -71.4749
0.25799 | -71.4194
0.25876 | -71.4306
0.25954 | -71.5137
0.26031 | -71.6122
0.26109 | -71.7716
0.26186 | -71.975
0.26264 | -72.2096
0.26341 | -72.4587
0.26419 | -72.7407
0.26496 | -73.007
0.26573 | -73.2453
0.26651 | -73.4028
0.26728 | -73.4975
0.26806 | -73.3966
0.26883 | -73.3697
0.26961 | -73.3039
0.27038 | -73.1238
0.27116 | -72.9059
0.27193 | -72.7029
0.27271 | -72.4941
0.27348 | -72.3175
0.27426 | -72.1478
0.27503 | -72.0166
0.27581 | -71.9394
0.27658 | -71.9075
0.27736 | -71.9014
0.27813 | -71.9559
0.27891 | -72.0568
0.27968 | -72.2082
0.28046 | -72.4142
0.28123 | -72.6365
0.282 | -72.8915
0.28278 | -73.1595
0.28355 | -73.4559
0.28433 | -73.7273
0.2851 | -73.9932
0.28588 | -74.2481
0.28665 | -74.4972
0.28743 | -74.6963
0.2882 | -74.8009
0.28898 | -74.8904
0.28975 | -74.989
0.29053 | -74.9149
0.2913 | -74.8937
0.29208 | -74.8592
0.29285 | -74.8095
0.29363 | -74.7137
0.2944 | -74.6131
0.29518 | -74.5356
0.29595 | -74.4579
0.29672 | -74.3963
0.2975 | -74.3687
0.29827 | -74.3873
0.29905 | -74.465
0.29982 | -74.5367
0.3006 | -74.6438
0.30137 | -74.7957
0.30215 | -74.9684
0.30292 | -75.1697
0.3037 | -75.3964
0.30447 | -75.646
0.30525 | -75.887
0.30602 | -76.0696
0.3068 | -76.3059
0.30757 | -76.4727
0.30834 | -76.5807
0.30912 | -76.6663
0.30989 | -76.7096
0.31067 | -76.6715
0.31144 | -76.5665
0.31222 | -76.3593
0.31299 | -76.2361
0.31377 | -76.0963
0.31454 | -75.9529
0.31532 | -75.7915
0.31609 | -75.6368
0.31687 | -75.5056
0.31764 | -75.4206
0.31842 | -75.3668
0.31919 | -75.3554
0.31997 | -75.403
0.32074 | -75.497
0.32152 | -75.6316
0.32229 | -75.8051
0.32306 | -76.0212
0.32384 | -76.2706
0.32461 | -76.5399
0.32539 | -76.841
0.32616 | -77.1458
0.32694 | -77.4279
0.32771 | -77.6986
0.32849 | -77.8771
0.32926 | -78.0474
0.33004 | -78.156
0.33081 | -78.236
0.33159 | -78.2771
0.33236 | -78.2254
0.33314 | -78.1215
0.33391 | -77.956
0.33469 | -77.6686
0.33546 | -77.4446
0.33624 | -77.2274
0.33701 | -77.0059
0.33779 | -76.7769
0.33856 | -76.5919
0.33933 | -76.4473
0.34011 | -76.3421
0.34088 | -76.2636
0.34166 | -76.2238
0.34243 | -76.2019
0.34321 | -76.2143
0.34398 | -76.2649
0.34476 | -76.3246
0.34553 | -76.4016
0.34631 | -76.481
0.34708 | -76.5586
0.34786 | -76.6656
0.34863 | -76.7898
0.34941 | -76.917
0.35018 | -77.0473
0.35096 | -77.1523
0.35173 | -77.2913
0.35251 | -77.4102
0.35328 | -77.5331
0.35406 | -77.6217
0.35483 | -77.6341
0.3556 | -77.644
0.35638 | -77.6308
0.35715 | -77.4493
0.35793 | -77.3843
0.3587 | -77.2893
0.35948 | -77.1803
0.36025 | -77.0599
0.36103 | -76.9287
0.3618 | -76.8225
0.36258 | -76.6975
0.36335 | -76.5422
0.36413 | -76.377
0.3649 | -76.2194
0.36568 | -76.0829
0.36645 | -75.9554
0.36723 | -75.8135
0.368 | -75.6842
0.36878 | -75.571
0.36955 | -75.4644
0.37032 | -75.3682
0.3711 | -75.2673
0.37187 | -75.1755
0.37265 | -75.1088
0.37342 | -75.0637
0.3742 | -75.011
0.37497 | -74.982
0.37575 | -74.9813
0.37652 | -74.993
0.3773 | -75.0413
0.37807 | -75.0991
0.37885 | -75.1887
0.37962 | -75.3761
0.3804 | -75.5276
0.38117 | -75.6802
0.38195 | -75.8489
0.38272 | -76.032
0.38349 | -76.235
0.38427 | -76.4796
0.38504 | -76.7059
0.38582 | -76.877
0.38659 | -77.0746
0.38737 | -77.2681
0.38814 | -77.4542
0.38892 | -77.6191
0.38969 | -77.75
0.39047 | -77.878
0.39124 | -78.002
0.39202 | -78.1466
0.39279 | -78.209
0.39357 | -78.2146
0.39434 | -78.2136
0.39512 | -78.2398
0.39589 | -78.2327
0.39667 | -78.1682
0.39744 | -78.109
0.39821 | -78.0489
0.39899 | -78.0136
0.39976 | -77.9596
0.40054 | -77.9155
0.40131 | -77.8775
0.40209 | -77.8724
0.40286 | -77.8658
0.40364 | -77.8632
0.40441 | -77.8619
0.40519 | -77.9016
0.40596 | -77.9536
0.40674 | -78.0059
0.40751 | -78.0553
0.40829 | -78.1145
0.40906 | -78.1857
0.40984 | -78.2685
0.41061 | -78.3527
0.41139 | -78.4588
0.41216 | -78.5735
0.41293 | -78.6836
0.41371 | -78.8124
0.41448 | -78.9461
0.41526 | -79.1116
0.41603 | -79.2524
0.41681 | -79.3509
0.41758 | -79.4632
0.41836 | -79.5908
0.41913 | -79.706
0.41991 | -79.8087
0.42068 | -79.9243
0.42146 | -79.9883
0.42223 | -80.0186
0.42301 | -80.0462
0.42378 | -80.0641
0.42456 | -79.9827
0.42533 | -79.9431
0.42611 | -79.9177
0.42688 | -79.9163
0.42766 | -79.8378
0.42843 | -79.7571
0.4292 | -79.6733
0.42998 | -79.6022
0.43075 | -79.5006
0.43153 | -79.4129
0.4323 | -79.3408
0.43308 | -79.2929
0.43385 | -79.2181
0.43463 | -79.1408
0.4354 | -79.0835
0.43618 | -79.0073
0.43695 | -78.9358
0.43773 | -78.8893
0.4385 | -78.8596
0.43928 | -78.8428
0.44005 | -78.8261
0.44083 | -78.8267
0.4416 | -78.8406
0.44238 | -78.8429
0.44315 | -78.8459
0.44393 | -78.823
0.4447 | -78.8031
0.44547 | -78.8178
0.44625 | -78.8258
0.44702 | -78.8328
0.4478 | -78.8372
0.44857 | -78.8594
0.44935 | -78.8668
0.45012 | -78.8847
0.4509 | -78.8962
0.45167 | -78.9068
0.45245 | -78.9462
0.45322 | -78.9784
0.454 | -79.0235
0.45477 | -79.0854
0.45555 | -79.148
0.45632 | -79.2212
0.45709 | -79.2916
0.45787 | -79.373
0.45865 | -79.4451
0.45942 | -79.5272
0.46019 | -79.627
0.46097 | -79.6982
0.46174 | -79.7802
0.46252 | -79.8642
0.46329 | -79.9486
0.46407 | -79.9978
0.46484 | -80.0343
0.46562 | -80.0799
0.46639 | -80.1229
0.46717 | -80.1554
0.46794 | -80.1716
0.46872 | -80.1514
0.46949 | -80.1635
0.47027 | -80.1628
0.47104 | -80.1495
0.47181 | -80.1312
0.47259 | -80.1146
0.47336 | -80.0994
0.47414 | -80.0558
0.47491 | -80.0117
0.47569 | -79.9722
0.47646 | -79.9418
0.47724 | -79.9399
0.47801 | -79.9069
0.47879 | -79.8379
0.47956 | -79.763
0.48034 | -79.7126
0.48111 | -79.6248
0.48189 | -79.527
0.48266 | -79.4298
0.48344 | -79.3346
0.48421 | -79.2383
0.48499 | -79.146
0.48576 | -79.051
0.48653 | -78.9435
0.48731 | -78.8415
0.48808 | -78.737
0.48886 | -78.6328
0.48963 | -78.5378
0.49041 | -78.4625
0.49118 | -78.4425
0.49196 | -78.4052
0.49273 | -78.3756
0.49351 | -78.362
0.49428 | -78.3577
0.49506 | -78.3732
0.49583 | -78.4059
0.49661 | -78.4513
0.49738 | -78.5132
0.49816 | -78.599
0.49893 | -78.6982
0.49971 | -78.8101
0.50048 | -78.9179
0.50126 | -79.0465
0.50203 | -79.1952
0.5028 | -79.356
0.50358 | -79.5239
0.50435 | -79.6805
0.50513 | -79.7964
0.5059 | -79.9126
0.50668 | -80.0223
0.50745 | -80.1211
0.50823 | -80.2042
0.509 | -80.2896
0.50978 | -80.3658
0.51055 | -80.4052
0.51133 | -80.4045
0.5121 | -80.3848
0.51288 | -80.3128
0.51365 | -80.1444
0.51443 | -80.0596
0.5152 | -79.9689
0.51597 | -79.8493
0.51675 | -79.6917
0.51753 | -79.5774
0.5183 | -79.4549
0.51907 | -79.3167
0.51985 | -79.1834
0.52062 | -79.0535
0.5214 | -78.9421
0.52217 | -78.8418
0.52295 | -78.7486
0.52372 | -78.6746
0.5245 | -78.6203
0.52527 | -78.5671
0.52605 | -78.5321
0.52682 | -78.5102
0.5276 | -78.4926
0.52837 | -78.5031
0.52915 | -78.5171
0.52992 | -78.5419
0.5307 | -78.5833
0.53147 | -78.6418
0.53224 | -78.6998
0.53302 | -78.7665
0.53379 | -78.8391
0.53457 | -78.916
0.53534 | -78.9981
0.53612 | -79.0758
0.53689 | -79.148
0.53767 | -79.2248
0.53844 | -79.3134
0.53922 | -79.3927
0.53999 | -79.4682
0.54077 | -79.5429
0.54154 | -79.6185
0.54232 | -79.7043
0.54309 | -79.8183
0.54387 | -79.9072
0.54464 | -79.9752
0.54541 | -80.0418
0.54619 | -80.1064
0.54697 | -80.1698
0.54774 | -80.2426
0.54851 | -80.331
0.54929 | -80.4107
0.55006 | -80.482
0.55084 | -80.5898
0.55161 | -80.6894
0.55239 | -80.7869
0.55316 | -80.8849
0.55394 | -80.9808
0.55471 | -81.1079
0.55549 | -81.2436
0.55626 | -81.3491
0.55704 | -81.4698
0.55781 | -81.5992
0.55859 | -81.7188
0.55936 | -81.8259
0.56014 | -81.916
0.56091 | -82.0007
0.56168 | -82.0926
0.56246 | -82.1999
0.56323 | -82.2878
0.56401 | -82.3904
0.56478 | -82.4891
0.56556 | -82.5716
0.56633 | -82.6299
0.56711 | -82.7077
0.56788 | -82.742
0.56866 | -82.7758
0.56943 | -82.8015
0.57021 | -82.7981
0.57098 | -82.7983
0.57176 | -82.8006
0.57253 | -82.7917
0.57331 | -82.7467
0.57408 | -82.7016
0.57486 | -82.6169
0.57563 | -82.5533
0.5764 | -82.5158
0.57718 | -82.4737
0.57795 | -82.4087
0.57873 | -82.326
0.5795 | -82.2422
0.58028 | -82.1713
0.58105 | -82.113
0.58183 | -82.0524
0.5826 | -81.9946
0.58338 | -81.9452
0.58415 | -81.9111
0.58493 | -81.8806
0.5857 | -81.8487
0.58648 | -81.8122
0.58725 | -81.7803
0.58803 | -81.7569
0.5888 | -81.7372
0.58957 | -81.7289
0.59035 | -81.7326
0.59113 | -81.7429
0.5919 | -81.7475
0.59267 | -81.7418
0.59345 | -81.7536
0.59422 | -81.7875
0.595 | -81.8287
0.59577 | -81.8728
0.59655 | -81.9176
0.59732 | -81.9719
0.5981 | -82.0278
0.59887 | -82.0931
0.59965 | -82.1617
0.60042 | -82.2261
0.6012 | -82.3042
0.60197 | -82.3897
0.60275 | -82.4732
0.60352 | -82.515
0.6043 | -82.5747
0.60507 | -82.6218
0.60584 | -82.663
0.60662 | -82.6916
0.60739 | -82.723
0.60817 | -82.7549
0.60894 | -82.7857
0.60972 | -82.7932
0.61049 | -82.8178
0.61127 | -82.8394
0.61204 | -82.8678
0.61282 | -82.8938
0.61359 | -82.908
0.61437 | -82.9103
0.61514 | -82.8928
0.61592 | -82.9095
0.61669 | -82.9129
0.61747 | -82.9435
0.61824 | -82.9482
0.61901 | -82.9263
0.61979 | -82.915
0.62057 | -82.8935
0.62134 | -82.8712
0.62211 | -82.8614
0.62289 | -82.8363
0.62366 | -82.8238
0.62444 | -82.8011
0.62521 | -82.7555
0.62599 | -82.7233
0.62676 | -82.7018
0.62754 | -82.6796
0.62831 | -82.6615
0.62909 | -82.6417
0.62986 | -82.6233
0.63064 | -82.6046
0.63141 | -82.5889
0.63219 | -82.5635
0.63296 | -82.5445
0.63374 | -82.524
0.63451 | -82.4883
0.63528 | -82.4536
0.63606 | -82.4276
0.63683 | -82.4069
0.63761 | -82.3577
0.63838 | -82.3124
0.63916 | -82.2585
0.63993 | -82.2176
0.64071 | -82.1783
0.64148 | -82.1372
0.64226 | -82.1025
0.64303 | -82.065
0.64381 | -82.0222
0.64458 | -81.9805
0.64536 | -81.9231
0.64613 | -81.852
0.64691 | -81.7918
0.64768 | -81.7309
0.64846 | -81.6765
0.64923 | -81.6279
0.65001 | -81.5766
0.65078 | -81.5328
0.65155 | -81.4918
0.65233 | -81.4429
0.6531 | -81.399
0.65388 | -81.3602
0.65465 | -81.326
0.65543 | -81.2961
0.6562 | -81.2696
0.65698 | -81.2452
0.65775 | -81.2327
0.65853 | -81.2268
0.6593 | -81.2325
0.66008 | -81.2459
0.66085 | -81.2685
0.66163 | -81.2721
0.6624 | -81.2934
0.66317 | -81.3239
0.66395 | -81.3601
0.66473 | -81.4095
0.6655 | -81.4598
0.66627 | -81.5257
0.66705 | -81.5886
0.66782 | -81.6399
0.6686 | -81.6963
0.66937 | -81.7503
0.67015 | -81.805
0.67092 | -81.8384
0.6717 | -81.891
0.67247 | -81.955
0.67325 | -82.0079
0.67402 | -82.0458
0.6748 | -82.0744
0.67557 | -82.1017
0.67635 | -82.1306
0.67712 | -82.154
0.6779 | -82.1739
0.67867 | -82.1894
0.67944 | -82.2063
0.68022 | -82.2177
0.68099 | -82.2244
0.68177 | -82.2475
0.68254 | -82.2376
0.68332 | -82.2104
0.68409 | -82.2019
0.68487 | -82.2051
0.68564 | -82.1943
0.68642 | -82.165
0.68719 | -82.0912
0.68797 | -82.0472
0.68874 | -82.0121
0.68952 | -81.9736
0.69029 | -81.9379
0.69107 | -81.9096
0.69184 | -81.8841
0.69261 | -81.8385
0.69339 | -81.7995
0.69417 | -81.7785
0.69494 | -81.7378
0.69571 | -81.6952
0.69649 | -81.6603
0.69726 | -81.6215
0.69804 | -81.5812
0.69881 | -81.5415
0.69959 | -81.5008
0.70036 | -81.4592
0.70114 | -81.4181
0.70191 | -81.3686
0.70269 | -81.3238
0.70346 | -81.2922
0.70424 | -81.2591
0.70501 | -81.2229
0.70579 | -81.1984
0.70656 | -81.1701
0.70734 | -81.1488
0.70811 | -81.1355
0.70888 | -81.1336
0.70966 | -81.1388
0.71043 | -81.1566
0.71121 | -81.169
0.71198 | -81.1746
0.71276 | -81.1949
0.71353 | -81.221
0.71431 | -81.2549
0.71508 | -81.3264
0.71586 | -81.3559
0.71663 | -81.405
0.71741 | -81.4655
0.71818 | -81.5276
0.71896 | -81.5879
0.71973 | -81.6428
0.72051 | -81.7148
0.72128 | -81.7896
0.72206 | -81.8527
0.72283 | -81.9145
0.72361 | -81.9733
0.72438 | -82.0355
0.72515 | -82.101
0.72593 | -82.152
0.7267 | -82.2112
0.72748 | -82.2715
0.72825 | -82.3284
0.72903 | -82.3698
0.7298 | -82.4037
0.73058 | -82.4353
0.73135 | -82.4577
0.73213 | -82.4677
0.7329 | -82.4997
0.73368 | -82.5314
0.73445 | -82.5258
0.73523 | -82.5051
0.736 | -82.4822
0.73677 | -82.4826
0.73755 | -82.4377
0.73833 | -82.4521
0.7391 | -82.4273
0.73987 | -82.3945
0.74065 | -82.3516
0.74142 | -82.3153
0.7422 | -82.271
0.74297 | -82.2292
0.74375 | -82.1964
0.74452 | -82.1768
0.7453 | -82.1634
0.74607 | -82.1394
0.74685 | -82.1168
0.74762 | -82.0833
0.7484 | -82.0587
0.74917 | -82.047
0.74995 | -82.0399
0.75072 | -82.0396
0.7515 | -82.0358
0.75227 | -82.0508
0.75304 | -82.0764
0.75382 | -82.1117
0.75459 | -82.1406
0.75537 | -82.1706
0.75614 | -82.2041
0.75692 | -82.2386
0.75769 | -82.2762
0.75847 | -82.3275
0.75924 | -82.388
0.76002 | -82.4425
0.76079 | -82.4915
0.76157 | -82.5467
0.76234 | -82.6008
0.76312 | -82.658
0.76389 | -82.7251
0.76467 | -82.7931
0.76544 | -82.8542
0.76621 | -82.9184
0.76699 | -82.98
0.76777 | -83.0346
0.76854 | -83.0922
0.76931 | -83.164
0.77009 | -83.2236
0.77086 | -83.276
0.77164 | -83.3254
0.77241 | -83.3819
0.77319 | -83.4246
0.77396 | -83.449
0.77474 | -83.4911
0.77551 | -83.5228
0.77629 | -83.5452
0.77706 | -83.5664
0.77784 | -83.5869
0.77861 | -83.6097
0.77939 | -83.6401
0.78016 | -83.6459
0.78094 | -83.6515
0.78171 | -83.6426
0.78248 | -83.6423
0.78326 | -83.6337
0.78403 | -83.6274
0.78481 | -83.6181
0.78558 | -83.619
0.78636 | -83.6171
0.78713 | -83.591
0.78791 | -83.5702
0.78868 | -83.5644
0.78946 | -83.5587
0.79023 | -83.5419
0.79101 | -83.521
0.79178 | -83.5038
0.79256 | -83.4853
0.79333 | -83.4545
0.79411 | -83.4262
0.79488 | -83.4035
0.79566 | -83.393
0.79643 | -83.3656
0.79721 | -83.3496
0.79798 | -83.3328
0.79875 | -83.3135
0.79953 | -83.296
0.8003 | -83.2767
0.80108 | -83.2609
0.80185 | -83.2524
0.80263 | -83.2479
0.8034 | -83.2453
0.80418 | -83.257
0.80495 | -83.2527
0.80573 | -83.2438
0.8065 | -83.2425
0.80728 | -83.2605
0.80805 | -83.266
0.80883 | -83.2706
0.8096 | -83.2834
0.81037 | -83.3045
0.81115 | -83.3231
0.81193 | -83.3524
0.8127 | -83.3903
0.81347 | -83.4261
0.81425 | -83.4585
0.81502 | -83.5014
0.8158 | -83.5465
0.81657 | -83.5814
0.81735 | -83.6278
0.81812 | -83.6775
0.8189 | -83.7194
0.81967 | -83.7615
0.82045 | -83.8109
0.82122 | -83.8649
0.822 | -83.9183
0.82277 | -83.9732
0.82355 | -84.0299
0.82432 | -84.0859
0.8251 | -84.1535
0.82587 | -84.2074
0.82664 | -84.2566
0.82742 | -84.3026
0.82819 | -84.3562
0.82897 | -84.4114
0.82974 | -84.4623
0.83052 | -84.5204
0.83129 | -84.5896
0.83207 | -84.6444
0.83284 | -84.6985
0.83362 | -84.751
0.83439 | -84.8004
0.83517 | -84.8419
0.83594 | -84.88
0.83672 | -84.9129
0.83749 | -84.9479
0.83827 | -84.9823
0.83904 | -85.0222
0.83981 | -85.06
0.84059 | -85.0987
0.84137 | -85.1526
0.84214 | -85.1852
0.84291 | -85.2044
0.84369 | -85.2213
0.84446 | -85.2503
0.84524 | -85.2892
0.84601 | -85.305
0.84679 | -85.3219
0.84756 | -85.3412
0.84834 | -85.3502
0.84911 | -85.3645
0.84989 | -85.3841
0.85066 | -85.4034
0.85144 | -85.4255
0.85221 | -85.4702
0.85299 | -85.5057
0.85376 | -85.5317
0.85454 | -85.5342
0.85531 | -85.5311
0.85608 | -85.5342
0.85686 | -85.5487
0.85763 | -85.5672
0.85841 | -85.5892
0.85918 | -85.5897
0.85996 | -85.5875
0.86073 | -85.5611
0.86151 | -85.5576
0.86228 | -85.5618
0.86306 | -85.5589
0.86383 | -85.5549
0.86461 | -85.5434
0.86538 | -85.5424
0.86616 | -85.5212
0.86693 | -85.4947
0.86771 | -85.4726
0.86848 | -85.4457
0.86925 | -85.4002
0.87003 | -85.3601
0.87081 | -85.2939
0.87158 | -85.1844
0.87235 | -85.1218
0.87313 | -85.0614
0.8739 | -85.0024
0.87468 | -84.9443
0.87545 | -84.8834
0.87623 | -84.8227
0.877 | -84.7659
0.87778 | -84.6946
0.87855 | -84.6204
0.87933 | -84.5514
0.8801 | -84.4843
0.88088 | -84.4112
0.88165 | -84.3396
0.88243 | -84.2783
0.8832 | -84.2169
0.88397 | -84.1605
0.88475 | -84.1161
0.88553 | -84.0628
0.8863 | -84.0134
0.88707 | -83.9671
0.88785 | -83.9273
0.88862 | -83.8873
0.8894 | -83.8471
0.89017 | -83.8152
0.89095 | -83.7831
0.89172 | -83.7594
0.8925 | -83.7291
0.89327 | -83.7066
0.89405 | -83.7162
0.89482 | -83.7012
0.8956 | -83.6922
0.89637 | -83.6902
0.89715 | -83.6895
0.89792 | -83.6877
0.8987 | -83.6935
0.89947 | -83.7033
0.90024 | -83.7102
0.90102 | -83.7143
0.90179 | -83.7244
0.90257 | -83.7383
0.90334 | -83.7572
0.90412 | -83.7838
0.90489 | -83.8102
0.90567 | -83.8326
0.90644 | -83.8575
0.90722 | -83.8844
0.90799 | -83.9131
0.90877 | -83.9433
0.90954 | -83.9686
0.91032 | -83.9997
0.91109 | -84.0411
0.91187 | -84.0587
0.91264 | -84.07
0.91341 | -84.0964
0.91419 | -84.1254
0.91497 | -84.1545
0.91574 | -84.1824
0.91651 | -84.2073
0.91729 | -84.2344
0.91806 | -84.2737
0.91884 | -84.32
0.91961 | -84.3493
0.92039 | -84.3733
0.92116 | -84.3949
0.92194 | -84.4078
0.92271 | -84.4376
0.92349 | -84.4571
0.92426 | -84.4873
0.92504 | -84.513
0.92581 | -84.5224
0.92659 | -84.5325
0.92736 | -84.5402
0.92814 | -84.5515
0.92891 | -84.5606
0.92968 | -84.5722
0.93046 | -84.5815
0.93123 | -84.5883
0.93201 | -84.5905
0.93278 | -84.5896
0.93356 | -84.5921
0.93433 | -84.616
0.93511 | -84.6482
0.93588 | -84.6495
0.93666 | -84.6583
0.93743 | -84.6616
0.93821 | -84.6791
0.93898 | -84.6789
0.93976 | -84.6777
0.94053 | -84.6846
0.94131 | -84.6912
0.94208 | -84.7043
0.94285 | -84.7216
0.94363 | -84.7421
0.94441 | -84.7696
0.94518 | -84.7914
0.94595 | -84.8093
0.94673 | -84.8327
0.9475 | -84.8512
0.94828 | -84.871
0.94905 | -84.8968
0.94983 | -84.9216
0.9506 | -84.938
0.95138 | -84.9472
0.95215 | -84.973
0.95293 | -84.9918
0.9537 | -85.016
0.95448 | -85.0486
0.95525 | -85.0788
0.95603 | -85.1147
0.9568 | -85.1471
0.95757 | -85.1718
0.95835 | -85.1978
0.95913 | -85.2234
0.9599 | -85.2403
0.96067 | -85.2597
0.96145 | -85.2845
0.96222 | -85.3021
0.963 | -85.3232
0.96377 | -85.3424
0.96455 | -85.363
0.96532 | -85.3769
0.9661 | -85.3641
0.96687 | -85.3746
0.96765 | -85.3956
0.96842 | -85.4069
0.9692 | -85.4167
0.96997 | -85.42
0.97075 | -85.421
0.97152 | -85.4254
0.9723 | -85.4318
0.97307 | -85.4325
0.97384 | -85.4341
0.97462 | -85.4205
0.97539 | -85.3923
0.97617 | -85.3815
0.97694 | -85.3808
0.97772 | -85.3654
0.97849 | -85.3564
0.97927 | -85.3434
0.98004 | -85.3193
0.98082 | -85.2966
0.98159 | -85.2687
0.98237 | -85.2394
0.98314 | -85.2205
0.98392 | -85.2036
0.98469 | -85.1877
0.98547 | -85.1581
0.98624 | -85.1392
0.98701 | -85.122
0.98779 | -85.1074
0.98857 | -85.0883
0.98934 | -85.0686
0.99011 | -85.0532
0.99089 | -85.0308
0.99166 | -85.0235
0.99244 | -85.0088
0.99321 | -84.9886
0.99399 | -84.9657
0.99476 | -84.9527
0.99554 | -84.9425
0.99631 | -84.9326
0.99709 | -84.9241
0.99786 | -84.9191
0.99864 | -84.9156
0.99941 | -84.9157
1.0002 | -84.9276
1.001 | -84.9321
1.0017 | -84.9248
1.0025 | -84.9245
1.0033 | -84.9232
1.0041 | -84.9299
1.0048 | -84.94
1.0056 | -84.9616
1.0064 | -84.9714
1.0072 | -84.9781
1.0079 | -84.9873
1.0087 | -85.0076
1.0095 | -85.0305
1.0103 | -85.05
1.011 | -85.0737
1.0118 | -85.0963
1.0126 | -85.1248
1.0134 | -85.1551
1.0141 | -85.1869
1.0149 | -85.2196
1.0157 | -85.2431
1.0165 | -85.2611
1.0172 | -85.2859
1.018 | -85.316
1.0188 | -85.3478
1.0196 | -85.3783
1.0203 | -85.4197
1.0211 | -85.4744
1.0219 | -85.5042
1.0227 | -85.5405
1.0234 | -85.5774
1.0242 | -85.6039
1.025 | -85.6272
1.0258 | -85.6513
1.0265 | -85.6768
1.0273 | -85.7102
1.0281 | -85.7149
1.0289 | -85.7412
1.0296 | -85.7559
1.0304 | -85.7803
1.0312 | -85.8023
1.0319 | -85.8148
1.0327 | -85.8308
1.0335 | -85.8471
1.0343 | -85.8522
1.035 | -85.8637
1.0358 | -85.8724
1.0366 | -85.8843
1.0374 | -85.9185
1.0381 | -85.9264
1.0389 | -85.9214
1.0397 | -85.911
1.0405 | -85.8966
1.0412 | -85.8913
1.042 | -85.8888
1.0428 | -85.8753
1.0436 | -85.8691
1.0443 | -85.845
1.0451 | -85.81
1.0459 | -85.8094
1.0467 | -85.8155
1.0474 | -85.8116
1.0482 | -85.7943
1.049 | -85.7733
1.0498 | -85.7603
1.0505 | -85.7366
1.0513 | -85.7316
1.0521 | -85.7124
1.0529 | -85.6929
1.0536 | -85.6753
1.0544 | -85.6538
1.0552 | -85.6196
1.056 | -85.5946
1.0567 | -85.5739
1.0575 | -85.5548
1.0583 | -85.5315
1.0591 | -85.5078
1.0598 | -85.4885
1.0606 | -85.4696
1.0614 | -85.4494
1.0622 | -85.4275
1.0629 | -85.4099
1.0637 | -85.4005
1.0645 | -85.3838
1.0653 | -85.3699
1.066 | -85.3523
1.0668 | -85.3426
1.0676 | -85.3257
1.0684 | -85.3054
1.0691 | -85.2825
1.0699 | -85.2669
1.0707 | -85.254
1.0715 | -85.2422
1.0722 | -85.2413
1.073 | -85.2274
1.0738 | -85.2141
1.0746 | -85.2089
1.0753 | -85.2078
1.0761 | -85.1948
1.0769 | -85.1761
1.0777 | -85.1551
1.0784 | -85.1431
1.0792 | -85.1419
1.08 | -85.1305
1.0808 | -85.1267
1.0815 | -85.1248
1.0823 | -85.1215
1.0831 | -85.1116
1.0839 | -85.0881
1.0846 | -85.0603
1.0854 | -85.0457
1.0862 | -85.0373
1.087 | -85.0293
1.0877 | -85.0218
1.0885 | -85.0154
1.0893 | -85.008
1.0901 | -85.009
1.0908 | -85.0112
1.0916 | -85.0051
1.0924 | -84.9985
1.0932 | -84.9833
1.0939 | -84.9668
1.0947 | -84.9513
1.0955 | -84.9382
1.0963 | -84.9198
1.097 | -84.9003
1.0978 | -84.8834
1.0986 | -84.8649
1.0994 | -84.8515
1.1001 | -84.8391
1.1009 | -84.8237
1.1017 | -84.8082
1.1025 | -84.7887
1.1032 | -84.7701
1.104 | -84.7445
1.1048 | -84.7276
1.1056 | -84.7108
1.1063 | -84.6901
1.1071 | -84.6676
1.1079 | -84.6477
1.1086 | -84.6275
1.1094 | -84.6116
1.1102 | -84.6
1.111 | -84.5877
1.1117 | -84.5763
1.1125 | -84.565
1.1133 | -84.5523
1.1141 | -84.5434
1.1148 | -84.5391
1.1156 | -84.5334
1.1164 | -84.5348
1.1172 | -84.5506
1.1179 | -84.5516
1.1187 | -84.5534
1.1195 | -84.5631
1.1203 | -84.5653
1.121 | -84.5821
1.1218 | -84.6033
1.1226 | -84.6096
1.1234 | -84.6195
1.1241 | -84.6333
1.1249 | -84.6499
1.1257 | -84.6681
1.1265 | -84.6934
1.1272 | -84.7197
1.128 | -84.7445
1.1288 | -84.7737
1.1296 | -84.7958
1.1303 | -84.8167
1.1311 | -84.85
1.1319 | -84.888
1.1327 | -84.9338
1.1334 | -84.9712
1.1342 | -84.9943
1.135 | -85.0216
1.1358 | -85.0655
1.1365 | -85.1068
1.1373 | -85.1448
1.1381 | -85.1854
1.1389 | -85.2245
1.1396 | -85.2517
1.1404 | -85.2774
1.1412 | -85.2974
1.142 | -85.3212
1.1427 | -85.345
1.1435 | -85.3733
1.1443 | -85.4007
1.1451 | -85.426
1.1458 | -85.4511
1.1466 | -85.4786
1.1474 | -85.5052
1.1482 | -85.5247
1.1489 | -85.5379
1.1497 | -85.5425
1.1505 | -85.5514
1.1513 | -85.5589
1.152 | -85.562
1.1528 | -85.5657
1.1536 | -85.569
1.1544 | -85.561
1.1551 | -85.5599
1.1559 | -85.563
1.1567 | -85.5606
1.1575 | -85.5485
1.1582 | -85.5319
1.159 | -85.5164
1.1598 | -85.4941
1.1606 | -85.4656
1.1613 | -85.4391
1.1621 | -85.4031
1.1629 | -85.3833
1.1637 | -85.3703
1.1644 | -85.3471
1.1652 | -85.29
1.166 | -85.2375
1.1668 | -85.1964
1.1675 | -85.151
1.1683 | -85.1046
1.1691 | -85.0545
1.1699 | -84.9976
1.1706 | -84.9449
1.1714 | -84.8933
1.1722 | -84.8461
1.173 | -84.8098
1.1737 | -84.7729
1.1745 | -84.7357
1.1753 | -84.6977
1.1761 | -84.6604
1.1768 | -84.6269
1.1776 | -84.594
1.1784 | -84.5301
1.1791 | -84.4677
1.1799 | -84.4102
1.1807 | -84.358
1.1815 | -84.3114
1.1822 | -84.2626
1.183 | -84.213
1.1838 | -84.1683
1.1846 | -84.1074
1.1853 | -84.0487
1.1861 | -83.9847
1.1869 | -83.9311
1.1877 | -83.8737
1.1884 | -83.8094
1.1892 | -83.7477
1.19 | -83.684
1.1908 | -83.6205
1.1915 | -83.564
1.1923 | -83.5063
1.1931 | -83.4479
1.1939 | -83.3896
1.1946 | -83.3293
1.1954 | -83.2582
1.1962 | -83.189
1.197 | -83.127
1.1977 | -83.0596
1.1985 | -82.9916
1.1993 | -82.9234
1.2001 | -82.8608
1.2008 | -82.8031
1.2016 | -82.7471
1.2024 | -82.6928
1.2032 | -82.6354
1.2039 | -82.5757
1.2047 | -82.5248
1.2055 | -82.4579
1.2063 | -82.3931
1.207 | -82.3446
1.2078 | -82.2926
1.2086 | -82.2394
1.2094 | -82.184
1.2101 | -82.1371
1.2109 | -82.0934
1.2117 | -82.0481
1.2125 | -82.0095
1.2132 | -81.9743
1.214 | -81.9354
1.2148 | -81.9
1.2156 | -81.867
1.2163 | -81.835
1.2171 | -81.8085
1.2179 | -81.7787
1.2187 | -81.7485
1.2194 | -81.7183
1.2202 | -81.6854
1.221 | -81.6606
1.2218 | -81.6416
1.2225 | -81.6177
1.2233 | -81.6008
1.2241 | -81.5924
1.2249 | -81.5833
1.2256 | -81.5686
1.2264 | -81.5591
1.2272 | -81.5488
1.228 | -81.5347
1.2287 | -81.5403
1.2295 | -81.5341
1.2303 | -81.5262
1.2311 | -81.5212
1.2318 | -81.5176
1.2326 | -81.5118
1.2334 | -81.51
1.2342 | -81.5143
1.2349 | -81.5186
1.2357 | -81.5207
1.2365 | -81.523
1.2373 | -81.5209
1.238 | -81.5197
1.2388 | -81.5253
1.2396 | -81.5303
1.2404 | -81.5376
1.2411 | -81.547
1.2419 | -81.5574
1.2427 | -81.5687
1.2435 | -81.579
1.2442 | -81.5824
1.245 | -81.594
1.2458 | -81.6087
1.2466 | -81.6239
1.2473 | -81.6398
1.2481 | -81.6557
1.2489 | -81.6742
1.2497 | -81.6945
1.2504 | -81.7184
1.2512 | -81.7282
1.252 | -81.7511
1.2528 | -81.7694
1.2535 | -81.7892
1.2543 | -81.8087
1.2551 | -81.824
1.2558 | -81.8458
1.2566 | -81.8671
1.2574 | -81.8876
1.2582 | -81.9033
1.2589 | -81.9154
1.2597 | -81.9293
1.2605 | -81.9413
1.2613 | -81.953
1.262 | -81.9666
1.2628 | -81.9776
1.2636 | -81.9889
1.2644 | -82.0184
1.2651 | -82.047
1.2659 | -82.0666
1.2667 | -82.0838
1.2675 | -82.1005
1.2682 | -82.12
1.269 | -82.1412
1.2698 | -82.1608
1.2706 | -82.1805
1.2713 | -82.201
1.2721 | -82.2281
1.2729 | -82.2538
1.2737 | -82.2803
1.2744 | -82.3024
1.2752 | -82.327
1.276 | -82.3558
1.2768 | -82.3826
1.2775 | -82.4081
1.2783 | -82.4359
1.2791 | -82.4578
1.2799 | -82.4804
1.2806 | -82.5074
1.2814 | -82.5365
1.2822 | -82.5692
1.283 | -82.6077
1.2837 | -82.6462
1.2845 | -82.6828
1.2853 | -82.7214
1.2861 | -82.7606
1.2868 | -82.8014
1.2876 | -82.8363
1.2884 | -82.871
1.2892 | -82.908
1.2899 | -82.9437
1.2907 | -82.9858
1.2915 | -83.0361
1.2923 | -83.0714
1.293 | -83.1152
1.2938 | -83.1673
1.2946 | -83.2118
1.2954 | -83.2476
1.2961 | -83.2896
1.2969 | -83.3341
1.2977 | -83.3809
1.2985 | -83.4332
1.2992 | -83.4771
1.3 | -83.5173
1.3008 | -83.5639
1.3016 | -83.6082
1.3023 | -83.6522
1.3031 | -83.6998
1.3039 | -83.7447
1.3047 | -83.7853
1.3054 | -83.8265
1.3062 | -83.8736
1.307 | -83.9273
1.3078 | -83.9785
1.3085 | -84.0292
1.3093 | -84.0827
1.3101 | -84.12
1.3109 | -84.1556
1.3116 | -84.2009
1.3124 | -84.2317
1.3132 | -84.2583
1.314 | -84.2889
1.3147 | -84.3183
1.3155 | -84.3352
1.3163 | -84.3467
1.3171 | -84.3707
1.3178 | -84.3888
1.3186 | -84.3671
1.3194 | -84.3756
1.3202 | -84.38
1.3209 | -84.3866
1.3217 | -84.3917
1.3225 | -84.3915
1.3233 | -84.3906
1.324 | -84.3878
1.3248 | -84.3778
1.3256 | -84.358
1.3263 | -84.3376
1.3271 | -84.3185
1.3279 | -84.2967
1.3287 | -84.2687
1.3294 | -84.2306
1.3302 | -84.1912
1.331 | -84.153
1.3318 | -84.1156
1.3325 | -84.0712
1.3333 | -84.0094
1.3341 | -83.9636
1.3349 | -83.933
1.3356 | -83.8912
1.3364 | -83.832
1.3372 | -83.777
1.338 | -83.7281
1.3387 | -83.6767
1.3395 | -83.6201
1.3403 | -83.5651
1.3411 | -83.5019
1.3418 | -83.4552
1.3426 | -83.4102
1.3434 | -83.3436
1.3442 | -83.2655
1.3449 | -83.203
1.3457 | -83.145
1.3465 | -83.0843
1.3473 | -83.0309
1.348 | -82.9791
1.3488 | -82.9204
1.3496 | -82.8542
1.3504 | -82.8013
1.3511 | -82.7505
1.3519 | -82.6949
1.3527 | -82.6429
1.3535 | -82.5898
1.3542 | -82.5283
1.355 | -82.4734
1.3558 | -82.4253
1.3566 | -82.3726
1.3573 | -82.3157
1.3581 | -82.2676
1.3589 | -82.2219
1.3597 | -82.1756
1.3604 | -82.1276
1.3612 | -82.0798
1.362 | -82.0342
1.3628 | -82.0082
1.3635 | -81.9685
1.3643 | -81.9312
1.3651 | -81.8958
1.3659 | -81.8632
1.3666 | -81.8333
1.3674 | -81.8044
1.3682 | -81.7765
1.369 | -81.7475
1.3697 | -81.7204
1.3705 | -81.6979
1.3713 | -81.6796
1.3721 | -81.6587
1.3728 | -81.6422
1.3736 | -81.6277
1.3744 | -81.6118
1.3752 | -81.5989
1.3759 | -81.5869
1.3767 | -81.5809
1.3775 | -81.5771
1.3783 | -81.5735
1.379 | -81.5734
1.3798 | -81.5752
1.3806 | -81.5782
1.3814 | -81.5851
1.3821 | -81.5934
1.3829 | -81.6036
1.3837 | -81.6189
1.3845 | -81.6352
1.3852 | -81.6684
1.386 | -81.6824
1.3868 | -81.7066
1.3876 | -81.7343
1.3883 | -81.7476
1.3891 | -81.7687
1.3899 | -81.8003
1.3907 | -81.831
1.3914 | -81.8533
1.3922 | -81.878
1.393 | -81.9092
1.3938 | -81.9419
1.3945 | -81.9775
1.3953 | -82.0149
1.3961 | -82.0461
1.3969 | -82.0818
1.3976 | -82.1221
1.3984 | -82.162
1.3992 | -82.2036
1.4 | -82.2469
1.4007 | -82.2892
1.4015 | -82.3301
1.4023 | -82.3711
1.403 | -82.4038
1.4038 | -82.4383
1.4046 | -82.4752
1.4054 | -82.5127
1.4061 | -82.5529
1.4069 | -82.5946
1.4077 | -82.6261
1.4085 | -82.671
1.4092 | -82.7134
1.41 | -82.7522
1.4108 | -82.7831
1.4116 | -82.8168
1.4123 | -82.8538
1.4131 | -82.8856
1.4139 | -82.9162
1.4147 | -82.9453
1.4154 | -82.9699
1.4162 | -82.993
1.417 | -83.0177
1.4178 | -83.0347
1.4185 | -83.0448
1.4193 | -83.0598
1.4201 | -83.0763
1.4209 | -83.0888
1.4216 | -83.0992
1.4224 | -83.1119
1.4232 | -83.1253
1.424 | -83.126
1.4247 | -83.1277
1.4255 | -83.1302
1.4263 | -83.1299
1.4271 | -83.126
1.4278 | -83.1152
1.4286 | -83.0967
1.4294 | -83.0777
1.4302 | -83.0244
1.4309 | -83.0153
1.4317 | -82.997
1.4325 | -82.9745
1.4333 | -82.9447
1.434 | -82.9152
1.4348 | -82.8845
1.4356 | -82.8516
1.4364 | -82.8224
1.4371 | -82.7917
1.4379 | -82.7543
1.4387 | -82.7178
1.4395 | -82.6806
1.4402 | -82.6397
1.441 | -82.6
1.4418 | -82.5746
1.4426 | -82.5477
1.4433 | -82.5025
1.4441 | -82.4594
1.4449 | -82.426
1.4457 | -82.3788
1.4464 | -82.3241
1.4472 | -82.2796
1.448 | -82.2468
1.4488 | -82.2075
1.4495 | -82.1687
1.4503 | -82.1324
1.4511 | -82.0949
1.4519 | -82.0589
1.4526 | -82.0241
1.4534 | -81.9963
1.4542 | -81.9655
1.455 | -81.9288
1.4557 | -81.8934
1.4565 | -81.8593
1.4573 | -81.8208
1.4581 | -81.7843
1.4588 | -81.7517
1.4596 | -81.7205
1.4604 | -81.6873
1.4612 | -81.6522
1.4619 | -81.6213
1.4627 | -81.5924
1.4635 | -81.5647
1.4643 | -81.5384
1.465 | -81.5109
1.4658 | -81.4782
1.4666 | -81.4473
1.4674 | -81.4224
1.4681 | -81.3892
1.4689 | -81.3587
1.4697 | -81.3316
1.4705 | -81.3074
1.4712 | -81.2839
1.472 | -81.2654
1.4728 | -81.2481
1.4735 | -81.2305
1.4743 | -81.2022
1.4751 | -81.171
1.4759 | -81.1522
1.4766 | -81.1366
1.4774 | -81.1208
1.4782 | -81.1053
1.479 | -81.0891
1.4797 | -81.0755
1.4805 | -81.0603
1.4813 | -81.043
1.4821 | -81.0254
1.4828 | -81.0087
1.4836 | -80.994
1.4844 | -80.9788
1.4852 | -80.9649
1.4859 | -80.9528
1.4867 | -80.9442
1.4875 | -80.9353
1.4883 | -80.9279
1.489 | -80.9224
1.4898 | -80.9209
1.4906 | -80.9142
1.4914 | -80.9055
1.4921 | -80.9016
1.4929 | -80.8989
1.4937 | -80.8943
1.4945 | -80.8896
1.4952 | -80.8855
1.496 | -80.8817
1.4968 | -80.8796
1.4976 | -80.8803
1.4983 | -80.8823
1.4991 | -80.8833
1.4999 | -80.8807
1.5007 | -80.8784
1.5014 | -80.8767
1.5022 | -80.8784
1.503 | -80.883
1.5038 | -80.8871
1.5045 | -80.8899
1.5053 | -80.8987
1.5061 | -80.905
1.5069 | -80.9085
1.5076 | -80.9165
1.5084 | -80.9214
1.5092 | -80.9255
1.51 | -80.9303
1.5107 | -80.9314
1.5115 | -80.9342
1.5123 | -80.9381
1.5131 | -80.9441
1.5138 | -80.9574
1.5146 | -80.9697
1.5154 | -80.9809
1.5162 | -80.9935
1.5169 | -81.0067
1.5177 | -81.0216
1.5185 | -81.0357
1.5193 | -81.0625
1.52 | -81.0778
1.5208 | -81.0927
1.5216 | -81.1129
1.5224 | -81.1363
1.5231 | -81.1575
1.5239 | -81.1774
1.5247 | -81.1984
1.5255 | -81.2222
1.5262 | -81.2448
1.527 | -81.2711
1.5278 | -81.3042
1.5286 | -81.3365
1.5293 | -81.3643
1.5301 | -81.3906
1.5309 | -81.4186
1.5317 | -81.4504
1.5324 | -81.4844
1.5332 | -81.5174
1.534 | -81.5548
1.5348 | -81.5925
1.5355 | -81.6287
1.5363 | -81.6657
1.5371 | -81.7043
1.5379 | -81.7489
1.5386 | -81.7964
1.5394 | -81.8362
1.5402 | -81.8755
1.541 | -81.913
1.5417 | -81.9512
1.5425 | -81.9964
1.5433 | -82.0439
1.5441 | -82.0875
1.5448 | -82.1319
1.5456 | -82.1781
1.5464 | -82.2212
1.5472 | -82.2652
1.5479 | -82.31
1.5487 | -82.3521
1.5495 | -82.3923
1.5502 | -82.4293
1.551 | -82.4779
1.5518 | -82.5271
1.5526 | -82.5655
1.5533 | -82.5937
1.5541 | -82.6152
1.5549 | -82.6345
1.5557 | -82.6486
1.5564 | -82.6594
1.5572 | -82.6805
1.558 | -82.7025
1.5588 | -82.7204
1.5595 | -82.7341
1.5603 | -82.7292
1.5611 | -82.7376
1.5619 | -82.7506
1.5626 | -82.7596
1.5634 | -82.7639
1.5642 | -82.7167
1.565 | -82.7129
1.5657 | -82.7063
1.5665 | -82.6907
1.5673 | -82.6747
1.5681 | -82.6699
1.5688 | -82.6549
1.5696 | -82.6365
1.5704 | -82.6158
1.5712 | -82.5889
1.5719 | -82.5647
1.5727 | -82.5358
1.5735 | -82.5125
1.5743 | -82.4827
1.575 | -82.4582
1.5758 | -82.4321
1.5766 | -82.3993
1.5774 | -82.3577
1.5781 | -82.3151
1.5789 | -82.2615
1.5797 | -82.1861
1.5805 | -82.1222
1.5812 | -82.0875
1.582 | -82.0544
1.5828 | -82.0076
1.5836 | -81.9552
1.5843 | -81.9011
1.5851 | -81.8364
1.5859 | -81.7738
1.5867 | -81.6941
1.5874 | -81.6418
1.5882 | -81.5847
1.589 | -81.5287
1.5898 | -81.4803
1.5905 | -81.4323
1.5913 | -81.3816
1.5921 | -81.3329
1.5929 | -81.2842
1.5936 | -81.23
1.5944 | -81.178
1.5952 | -81.1319
1.596 | -81.0835
1.5967 | -81.0244
1.5975 | -80.9767
1.5983 | -80.9241
1.5991 | -80.8727
1.5998 | -80.8255
1.6006 | -80.7643
1.6014 | -80.7068
1.6022 | -80.6638
1.6029 | -80.6211
1.6037 | -80.579
1.6045 | -80.5371
1.6053 | -80.4936
1.606 | -80.4537
1.6068 | -80.4186
1.6076 | -80.3731
1.6084 | -80.3344
1.6091 | -80.3156
1.6099 | -80.2795
1.6107 | -80.2489
1.6115 | -80.2203
1.6122 | -80.1921
1.613 | -80.1646
1.6138 | -80.1367
1.6146 | -80.112
1.6153 | -80.0885
1.6161 | -80.068
1.6169 | -80.0505
1.6177 | -80.0314
1.6184 | -80.0141
1.6192 | -79.9986
1.62 | -79.9824
1.6208 | -79.9678
1.6215 | -79.9548
1.6223 | -79.944
1.6231 | -79.9345
1.6239 | -79.9256
1.6246 | -79.9172
1.6254 | -79.908
1.6262 | -79.8848
1.6269 | -79.8729
1.6277 | -79.8681
1.6285 | -79.8616
1.6293 | -79.8562
1.63 | -79.8557
1.6308 | -79.8554
1.6316 | -79.87
1.6324 | -79.8727
1.6331 | -79.876
1.6339 | -79.8806
1.6347 | -79.884
1.6355 | -79.8919
1.6362 | -79.8971
1.637 | -79.9057
1.6378 | -79.9175
1.6386 | -79.9286
1.6393 | -79.9411
1.6401 | -79.9549
1.6409 | -79.971
1.6417 | -79.9906
1.6424 | -80.0085
1.6432 | -80.0259
1.644 | -80.0454
1.6448 | -80.0657
1.6455 | -80.081
1.6463 | -80.0966
1.6471 | -80.1177
1.6479 | -80.134
1.6486 | -80.1505
1.6494 | -80.1715
1.6502 | -80.1956
1.651 | -80.2223
1.6517 | -80.249
1.6525 | -80.2772
1.6533 | -80.3069
1.6541 | -80.3366
1.6548 | -80.3624
1.6556 | -80.3866
1.6564 | -80.4062
1.6572 | -80.4279
1.6579 | -80.4504
1.6587 | -80.4733
1.6595 | -80.5015
1.6603 | -80.5234
1.661 | -80.5398
1.6618 | -80.5579
1.6626 | -80.5723
1.6634 | -80.592
1.6641 | -80.6168
1.6649 | -80.6404
1.6657 | -80.6645
1.6665 | -80.6926
1.6672 | -80.721
1.668 | -80.7441
1.6688 | -80.7699
1.6696 | -80.7966
1.6703 | -80.8165
1.6711 | -80.8362
1.6719 | -80.8563
1.6727 | -80.8674
1.6734 | -80.8884
1.6742 | -80.9022
1.675 | -80.9143
1.6758 | -80.9134
1.6765 | -80.9238
1.6773 | -80.936
1.6781 | -80.9479
1.6789 | -80.9599
1.6796 | -80.9657
1.6804 | -80.9645
1.6812 | -80.9711
1.682 | -80.9818
1.6827 | -80.9764
1.6835 | -80.97
1.6843 | -80.9673
1.6851 | -80.9674
1.6858 | -80.963
1.6866 | -80.9581
1.6874 | -80.9586
1.6882 | -80.9518
1.6889 | -80.9431
1.6897 | -80.9316
1.6905 | -80.9222
1.6913 | -80.9106
1.692 | -80.9008
1.6928 | -80.8888
1.6936 | -80.8832
1.6944 | -80.8903
1.6951 | -80.8796
1.6959 | -80.8424
1.6967 | -80.8056
1.6974 | -80.7803
1.6982 | -80.7534
1.699 | -80.7348
1.6998 | -80.7115
1.7005 | -80.6948
1.7013 | -80.6745
1.7021 | -80.6518
1.7029 | -80.6382
1.7036 | -80.6142
1.7044 | -80.5932
1.7052 | -80.5732
1.706 | -80.5496
1.7067 | -80.5258
1.7075 | -80.5027
1.7083 | -80.4766
1.7091 | -80.4545
1.7098 | -80.4361
1.7106 | -80.416
1.7114 | -80.3824
1.7122 | -80.3537
1.7129 | -80.3384
1.7137 | -80.3108
1.7145 | -80.2912
1.7153 | -80.2743
1.716 | -80.2548
1.7168 | -80.2417
1.7176 | -80.2239
1.7184 | -80.2008
1.7191 | -80.1803
1.7199 | -80.1594
1.7207 | -80.1487
1.7215 | -80.1312
1.7222 | -80.1166
1.723 | -80.1018
1.7238 | -80.0794
1.7246 | -80.0606
1.7253 | -80.0462
1.7261 | -80.0272
1.7269 | -80.011
1.7277 | -80.0004
1.7284 | -79.9872
1.7292 | -79.9732
1.73 | -79.9598
1.7308 | -79.9484
1.7315 | -79.9366
1.7323 | -79.9234
1.7331 | -79.9151
1.7339 | -79.9072
1.7346 | -79.8991
1.7354 | -79.8883
1.7362 | -79.8826
1.737 | -79.8778
1.7377 | -79.8743
1.7385 | -79.8736
1.7393 | -79.8714
1.7401 | -79.8704
1.7408 | -79.8697
1.7416 | -79.8678
1.7424 | -79.861
1.7432 | -79.855
1.7439 | -79.8425
1.7447 | -79.8511
1.7455 | -79.8541
1.7463 | -79.8525
1.747 | -79.8533
1.7478 | -79.8569
1.7486 | -79.858
1.7494 | -79.8596
1.7501 | -79.8636
1.7509 | -79.865
1.7517 | -79.8628
1.7525 | -79.8624
1.7532 | -79.8644
1.754 | -79.8713
1.7548 | -79.8766
1.7556 | -79.8849
1.7563 | -79.8905
1.7571 | -79.8956
1.7579 | -79.8974
1.7587 | -79.9024
1.7594 | -79.9102
1.7602 | -79.9201
1.761 | -79.9265
1.7618 | -79.9297
1.7625 | -79.9383
1.7633 | -79.9479
1.7641 | -79.9569
1.7649 | -79.9664
1.7656 | -79.9763
1.7664 | -79.9909
1.7672 | -79.9988
1.768 | -80.0093
1.7687 | -80.0197
1.7695 | -80.0297
1.7703 | -80.038
1.7711 | -80.0487
1.7718 | -80.0589
1.7726 | -80.069
1.7734 | -80.0776
1.7741 | -80.0858
1.7749 | -80.0931
1.7757 | -80.1014
1.7765 | -80.1123
1.7772 | -80.1278
1.778 | -80.1338
1.7788 | -80.1341
1.7796 | -80.1382
1.7803 | -80.1449
1.7811 | -80.1527
1.7819 | -80.1621
1.7827 | -80.1695
1.7834 | -80.1748
1.7842 | -80.1671
1.785 | -80.1477
1.7858 | -80.1308
1.7865 | -80.1229
1.7873 | -80.1096
1.7881 | -80.0943
1.7889 | -80.0858
1.7896 | -80.0807
1.7904 | -80.0771
1.7912 | -80.0724
1.792 | -80.0647
1.7927 | -80.0482
1.7935 | -80.0322
1.7943 | -80.0231
1.7951 | -80.0166
1.7958 | -80.0071
1.7966 | -79.9952
1.7974 | -79.9808
1.7982 | -79.9641
1.7989 | -79.9478
1.7997 | -79.9303
1.8005 | -79.9103
1.8013 | -79.8891
1.802 | -79.8686
1.8028 | -79.8506
1.8036 | -79.8299
1.8044 | -79.8078
1.8051 | -79.7755
1.8059 | -79.7408
1.8067 | -79.7103
1.8075 | -79.6817
1.8082 | -79.659
1.809 | -79.6336
1.8098 | -79.6005
1.8106 | -79.5467
1.8113 | -79.5072
1.8121 | -79.4721
1.8129 | -79.4402
1.8137 | -79.4049
1.8144 | -79.3735
1.8152 | -79.3418
1.816 | -79.3104
1.8168 | -79.2778
1.8175 | -79.2434
1.8183 | -79.2149
1.8191 | -79.1928
1.8199 | -79.1703
1.8206 | -79.1386
1.8214 | -79.1014
1.8222 | -79.0597
1.823 | -79.0204
1.8237 | -78.9838
1.8245 | -78.9493
1.8253 | -78.9158
1.8261 | -78.8828
1.8268 | -78.8435
1.8276 | -78.8007
1.8284 | -78.7647
1.8292 | -78.7303
1.8299 | -78.696
1.8307 | -78.6587
1.8315 | -78.6236
1.8323 | -78.5938
1.833 | -78.5614
1.8338 | -78.5265
1.8346 | -78.4899
1.8354 | -78.4566
1.8361 | -78.4326
1.8369 | -78.4064
1.8377 | -78.3755
1.8385 | -78.3498
1.8392 | -78.3199
1.84 | -78.2871
1.8408 | -78.2603
1.8416 | -78.2336
1.8423 | -78.2084
1.8431 | -78.1852
1.8439 | -78.1593
1.8446 | -78.1324
1.8454 | -78.1122
1.8462 | -78.0871
1.847 | -78.0627
1.8477 | -78.0404
1.8485 | -78.0228
1.8493 | -78.0087
1.8501 | -77.9988
1.8508 | -77.9808
1.8516 | -77.9618
1.8524 | -77.9479
1.8532 | -77.9347
1.8539 | -77.9211
1.8547 | -77.918
1.8555 | -77.8954
1.8563 | -77.8785
1.857 | -77.8684
1.8578 | -77.8589
1.8586 | -77.8528
1.8594 | -77.8449
1.8601 | -77.8332
1.8609 | -77.8185
1.8617 | -77.806
1.8625 | -77.8001
1.8632 | -77.796
1.864 | -77.7927
1.8648 | -77.791
1.8656 | -77.789
1.8663 | -77.7858
1.8671 | -77.7811
1.8679 | -77.7799
1.8687 | -77.7852
1.8694 | -77.7884
1.8702 | -77.798
1.871 | -77.8081
1.8718 | -77.8129
1.8725 | -77.8144
1.8733 | -77.8172
1.8741 | -77.8198
1.8749 | -77.8221
1.8756 | -77.8265
1.8764 | -77.8319
1.8772 | -77.8421
1.878 | -77.8413
1.8787 | -77.8413
1.8795 | -77.8458
1.8803 | -77.8536
1.8811 | -77.8595
1.8818 | -77.8647
1.8826 | -77.8686
1.8834 | -77.873
1.8842 | -77.8789
1.8849 | -77.8839
1.8857 | -77.893
1.8865 | -77.9004
1.8873 | -77.906
1.888 | -77.9125
1.8888 | -77.9197
1.8896 | -77.9305
1.8904 | -77.9387
1.8911 | -77.9454
1.8919 | -77.9531
1.8927 | -77.9642
1.8935 | -77.9734
1.8942 | -77.9804
1.895 | -77.9877
1.8958 | -77.9961
1.8966 | -78.0053
1.8973 | -78.0129
1.8981 | -78.0178
1.8989 | -78.0232
1.8997 | -78.0302
1.9004 | -78.0385
1.9012 | -78.0452
1.902 | -78.0511
1.9028 | -78.0603
1.9035 | -78.0695
1.9043 | -78.0772
1.9051 | -78.0848
1.9059 | -78.0919
1.9066 | -78.1023
1.9074 | -78.1168
1.9082 | -78.1316
1.909 | -78.1457
1.9097 | -78.1556
1.9105 | -78.1665
1.9113 | -78.1792
1.9121 | -78.1924
1.9128 | -78.2038
1.9136 | -78.22
1.9144 | -78.2414
1.9152 | -78.2609
1.9159 | -78.274
1.9167 | -78.2898
1.9175 | -78.3037
1.9183 | -78.3145
1.919 | -78.3292
1.9198 | -78.3435
1.9206 | -78.3564
1.9213 | -78.3698
1.9221 | -78.3949
1.9229 | -78.411
1.9237 | -78.4325
1.9244 | -78.4541
1.9252 | -78.4745
1.926 | -78.4927
1.9268 | -78.5102
1.9275 | -78.5287
1.9283 | -78.55
1.9291 | -78.5743
1.9299 | -78.597
1.9306 | -78.6195
1.9314 | -78.6446
1.9322 | -78.6655
1.933 | -78.6884
1.9337 | -78.7153
1.9345 | -78.7404
1.9353 | -78.7637
1.9361 | -78.7999
1.9368 | -78.8355
1.9376 | -78.863
1.9384 | -78.8905
1.9392 | -78.915
1.9399 | -78.9375
1.9407 | -78.9669
1.9415 | -78.9994
1.9423 | -79.0339
1.943 | -79.0709
1.9438 | -79.102
1.9446 | -79.1361
1.9454 | -79.1666
1.9461 | -79.1985
1.9469 | -79.2284
1.9477 | -79.2584
1.9485 | -79.2878
1.9492 | -79.316
1.95 | -79.3446
1.9508 | -79.3739
1.9516 | -79.4062
1.9523 | -79.436
1.9531 | -79.4652
1.9539 | -79.4997
1.9547 | -79.5379
1.9554 | -79.5728
1.9562 | -79.6069
1.957 | -79.635
1.9578 | -79.662
1.9585 | -79.6888
1.9593 | -79.7225
1.9601 | -79.7633
1.9609 | -79.7951
1.9616 | -79.825
1.9624 | -79.8542
1.9632 | -79.8901
1.964 | -79.9246
1.9647 | -79.9538
1.9655 | -79.9868
1.9663 | -80.0195
1.9671 | -80.0429
1.9678 | -80.0771
1.9686 | -80.1115
1.9694 | -80.1451
1.9702 | -80.1724
1.9709 | -80.2
1.9717 | -80.2305
1.9725 | -80.2587
1.9733 | -80.2911
1.974 | -80.3264
1.9748 | -80.3537
1.9756 | -80.3852
1.9764 | -80.4159
1.9771 | -80.4412
1.9779 | -80.4702
1.9787 | -80.502
1.9795 | -80.5265
1.9802 | -80.5551
1.981 | -80.5824
1.9818 | -80.6082
1.9826 | -80.6292
1.9833 | -80.6508
1.9841 | -80.6721
1.9849 | -80.6972
1.9857 | -80.7224
1.9864 | -80.7428
1.9872 | -80.7645
1.988 | -80.7872
1.9888 | -80.7986
1.9895 | -80.8167
1.9903 | -80.8502
1.9911 | -80.8922
1.9918 | -80.9229
1.9926 | -80.942
1.9934 | -80.9638
1.9942 | -80.9821
1.9949 | -81.0001
1.9957 | -81.0256
1.9965 | -81.0586
1.9973 | -81.0926
1.998 | -81.1077
1.9988 | -81.118
1.9996 | -81.1313
2.0004 | -81.1453
2.0011 | -81.1601
2.0019 | -81.1792
2.0027 | -81.2011
2.0035 | -81.2181
2.0042 | -81.2316
2.005 | -81.2473
2.0058 | -81.2638
2.0066 | -81.2803
2.0073 | -81.2977
2.0081 | -81.3182
2.0089 | -81.3335
2.0097 | -81.3501
2.0104 | -81.3698
2.0112 | -81.3931
2.012 | -81.4096
2.0128 | -81.4269
2.0135 | -81.4463
2.0143 | -81.4703
2.0151 | -81.5024
2.0159 | -81.5327
2.0166 | -81.5594
2.0174 | -81.5767
2.0182 | -81.594
2.019 | -81.6133
2.0197 | -81.6371
2.0205 | -81.6615
2.0213 | -81.6841
2.0221 | -81.7101
2.0228 | -81.7315
2.0236 | -81.7522
2.0244 | -81.7697
2.0252 | -81.7799
2.0259 | -81.7954
2.0267 | -81.8152
2.0275 | -81.8355
2.0283 | -81.8579
2.029 | -81.8824
2.0298 | -81.9012
2.0306 | -81.9214
2.0314 | -81.9446
2.0321 | -81.9675
2.0329 | -81.9911
2.0337 | -82.0185
2.0345 | -82.0429
2.0352 | -82.0716
2.036 | -82.1008
2.0368 | -82.1309
2.0376 | -82.1651
2.0383 | -82.1899
2.0391 | -82.2138
2.0399 | -82.2422
2.0407 | -82.2732
2.0414 | -82.2999
2.0422 | -82.3307
2.043 | -82.3624
2.0438 | -82.3947
2.0445 | -82.4236
2.0453 | -82.4477
2.0461 | -82.4726
2.0469 | -82.498
2.0476 | -82.5222
2.0484 | -82.5477
2.0492 | -82.5609
2.05 | -82.5767
2.0507 | -82.6009
2.0515 | -82.634
2.0523 | -82.6621
2.0531 | -82.6854
2.0538 | -82.71
2.0546 | -82.734
2.0554 | -82.758
2.0562 | -82.778
2.0569 | -82.7992
2.0577 | -82.8218
2.0585 | -82.8385
2.0593 | -82.8563
2.06 | -82.87
2.0608 | -82.8889
2.0616 | -82.907
2.0624 | -82.9217
2.0631 | -82.9403
2.0639 | -82.9551
2.0647 | -82.9702
2.0654 | -82.9854
2.0662 | -82.9943
2.067 | -83.0043
2.0678 | -83.0192
2.0685 | -83.0311
2.0693 | -83.0466
2.0701 | -83.0598
2.0709 | -83.0727
2.0716 | -83.0886
2.0724 | -83.1149
2.0732 | -83.129
2.074 | -83.1336
2.0747 | -83.1446
2.0755 | -83.1564
2.0763 | -83.1591
2.0771 | -83.165
2.0778 | -83.1746
2.0786 | -83.168
2.0794 | -83.1671
2.0802 | -83.1754
2.0809 | -83.1874
2.0817 | -83.1938
2.0825 | -83.1984
2.0833 | -83.2003
2.084 | -83.1995
2.0848 | -83.2028
2.0856 | -83.2016
2.0864 | -83.1984
2.0871 | -83.1948
2.0879 | -83.1927
2.0887 | -83.1904
2.0895 | -83.1858
2.0902 | -83.1823
2.091 | -83.1821
2.0918 | -83.1829
2.0926 | -83.1732
2.0933 | -83.1568
2.0941 | -83.1503
2.0949 | -83.1476
2.0957 | -83.139
2.0964 | -83.1338
2.0972 | -83.1188
2.098 | -83.1061
2.0988 | -83.1056
2.0995 | -83.1072
2.1003 | -83.1085
2.1011 | -83.0944
2.1019 | -83.0751
2.1026 | -83.0529
2.1034 | -83.0355
2.1042 | -83.0231
2.105 | -83.014
2.1057 | -82.9993
2.1065 | -82.9817
2.1073 | -82.9648
2.1081 | -82.9503
2.1088 | -82.9358
2.1096 | -82.9215
2.1104 | -82.9092
2.1112 | -82.8893
2.1119 | -82.8736
2.1127 | -82.8572
2.1135 | -82.845
2.1143 | -82.8328
2.115 | -82.8187
2.1158 | -82.8042
2.1166 | -82.7865
2.1174 | -82.7696
2.1181 | -82.7518
2.1189 | -82.7355
2.1197 | -82.7227
2.1205 | -82.7088
2.1212 | -82.6917
2.122 | -82.6739
2.1228 | -82.6539
2.1236 | -82.6407
2.1243 | -82.6288
2.1251 | -82.6125
2.1259 | -82.5963
2.1267 | -82.584
2.1274 | -82.5719
2.1282 | -82.5586
2.129 | -82.5493
2.1298 | -82.546
2.1305 | -82.5331
2.1313 | -82.5155
2.1321 | -82.501
2.1329 | -82.4916
2.1336 | -82.4837
2.1344 | -82.4744
2.1352 | -82.4642
2.136 | -82.4518
2.1367 | -82.4399
2.1375 | -82.4294
2.1383 | -82.4191
2.1391 | -82.4096
2.1398 | -82.3986
2.1406 | -82.3861
2.1414 | -82.3758
2.1421 | -82.3661
2.1429 | -82.3577
2.1437 | -82.3546
2.1445 | -82.3491
2.1452 | -82.343
2.146 | -82.3346
2.1468 | -82.3308
2.1476 | -82.3237
2.1483 | -82.3191
2.1491 | -82.3166
2.1499 | -82.3123
2.1507 | -82.307
2.1514 | -82.3023
2.1522 | -82.2974
2.153 | -82.2931
2.1538 | -82.2876
2.1545 | -82.2822
2.1553 | -82.2813
2.1561 | -82.2767
2.1569 | -82.2745
2.1576 | -82.2736
2.1584 | -82.2715
2.1592 | -82.2681
2.16 | -82.2676
2.1607 | -82.2674
2.1615 | -82.2676
2.1623 | -82.267
2.1631 | -82.2672
2.1638 | -82.2653
2.1646 | -82.269
2.1654 | -82.2707
2.1662 | -82.2704
2.1669 | -82.2657
2.1677 | -82.2744
2.1685 | -82.2714
2.1693 | -82.2685
2.17 | -82.2699
2.1708 | -82.2744
2.1716 | -82.2816
2.1724 | -82.2951
2.1731 | -82.3067
2.1739 | -82.3109
2.1747 | -82.3092
2.1755 | -82.3104
2.1762 | -82.3123
2.177 | -82.3134
2.1778 | -82.3166
2.1786 | -82.3206
2.1793 | -82.3249
2.1801 | -82.3265
2.1809 | -82.3287
2.1817 | -82.3303
2.1824 | -82.335
2.1832 | -82.3401
2.184 | -82.3424
2.1848 | -82.3463
2.1855 | -82.3496
2.1863 | -82.3424
2.1871 | -82.3322
2.1879 | -82.3342
2.1886 | -82.3371
2.1894 | -82.3398
2.1902 | -82.3422
2.191 | -82.346
2.1917 | -82.3501
2.1925 | -82.353
2.1933 | -82.3555
2.1941 | -82.3567
2.1948 | -82.3599
2.1956 | -82.3648
2.1964 | -82.3671
2.1972 | -82.3688
2.1979 | -82.3786
2.1987 | -82.3893
2.1995 | -82.3946
2.2003 | -82.3968
2.201 | -82.3977
2.2018 | -82.3997
2.2026 | -82.4012
2.2034 | -82.4044
2.2041 | -82.405
2.2049 | -82.4061
2.2057 | -82.4063
2.2065 | -82.4087
2.2072 | -82.4019
2.208 | -82.3954
2.2088 | -82.3953
2.2096 | -82.4002
2.2103 | -82.4003
2.2111 | -82.4022
2.2119 | -82.4023
2.2127 | -82.3969
2.2134 | -82.3936
2.2142 | -82.3862
2.215 | -82.3746
2.2157 | -82.3677
2.2165 | -82.3643
2.2173 | -82.3569
2.2181 | -82.3442
2.2188 | -82.3349
2.2196 | -82.3273
2.2204 | -82.3218
2.2212 | -82.3138
2.2219 | -82.2924
2.2227 | -82.2823
2.2235 | -82.2761
2.2243 | -82.2702
2.225 | -82.2605
2.2258 | -82.2513
2.2266 | -82.2451
2.2274 | -82.2386
2.2281 | -82.2307
2.2289 | -82.2212
2.2297 | -82.2106
2.2305 | -82.1953
2.2312 | -82.1768
2.232 | -82.1669
2.2328 | -82.1559
2.2336 | -82.1403
2.2343 | -82.1277
2.2351 | -82.1125
2.2359 | -82.0908
2.2367 | -82.0699
2.2374 | -82.0508
2.2382 | -82.0392
2.239 | -82.0296
2.2398 | -82.0161
2.2405 | -81.9989
2.2413 | -81.982
2.2421 | -81.9639
2.2429 | -81.9446
2.2436 | -81.9259
2.2444 | -81.9115
2.2452 | -81.9011
2.246 | -81.8843
2.2467 | -81.8643
2.2475 | -81.8478
2.2483 | -81.8317
2.2491 | -81.8207
2.2498 | -81.8095
2.2506 | -81.7924
2.2514 | -81.7731
2.2522 | -81.7563
2.2529 | -81.7397
2.2537 | -81.7173
2.2545 | -81.6935
2.2553 | -81.6755
2.256 | -81.6608
2.2568 | -81.6445
2.2576 | -81.6253
2.2584 | -81.6045
2.2591 | -81.5873
2.2599 | -81.5724
2.2607 | -81.5581
2.2615 | -81.543
2.2622 | -81.5242
2.263 | -81.5039
2.2638 | -81.483
2.2646 | -81.4628
2.2653 | -81.4462
2.2661 | -81.4248
2.2669 | -81.4
2.2677 | -81.3796
2.2684 | -81.3609
2.2692 | -81.3373
2.27 | -81.3183
2.2708 | -81.3011
2.2715 | -81.285
2.2723 | -81.2695
2.2731 | -81.2527
2.2739 | -81.2338
2.2746 | -81.2155
2.2754 | -81.1995
2.2762 | -81.1857
2.277 | -81.1733
2.2777 | -81.1572
2.2785 | -81.143
2.2793 | -81.1319
2.2801 | -81.1219
2.2808 | -81.111
2.2816 | -81.0966
2.2824 | -81.0827
2.2832 | -81.0672
2.2839 | -81.0502
2.2847 | -81.0367
2.2855 | -81.0254
2.2862 | -81.0158
2.287 | -81.005
2.2878 | -80.9942
2.2886 | -80.9818
2.2893 | -80.9704
2.2901 | -80.9589
2.2909 | -80.9484
2.2917 | -80.9374
2.2924 | -80.9233
2.2932 | -80.9075
2.294 | -80.8955
2.2948 | -80.8876
2.2955 | -80.8821
2.2963 | -80.8727
2.2971 | -80.8614
2.2979 | -80.8508
2.2986 | -80.8419
2.2994 | -80.8332
2.3002 | -80.8258
2.301 | -80.8204
2.3017 | -80.8123
2.3025 | -80.8068
2.3033 | -80.7989
2.3041 | -80.7854
2.3048 | -80.7741
2.3056 | -80.7681
2.3064 | -80.7627
2.3072 | -80.7568
2.3079 | -80.7503
2.3087 | -80.7455
2.3095 | -80.7433
2.3103 | -80.7375
2.311 | -80.7346
2.3118 | -80.7338
2.3126 | -80.7317
2.3134 | -80.7275
2.3141 | -80.7237
2.3149 | -80.7187
2.3157 | -80.7135
2.3165 | -80.7083
2.3172 | -80.7066
2.318 | -80.7068
2.3188 | -80.7079
2.3196 | -80.7078
2.3203 | -80.7099
2.3211 | -80.7096
2.3219 | -80.7094
2.3227 | -80.7096
2.3234 | -80.7152
2.3242 | -80.7227
2.325 | -80.7247
2.3258 | -80.7272
2.3265 | -80.7333
2.3273 | -80.7392
2.3281 | -80.7417
2.3289 | -80.7441
2.3296 | -80.7469
2.3304 | -80.7487
2.3312 | -80.7507
2.332 | -80.7547
2.3327 | -80.7606
2.3335 | -80.7665
2.3343 | -80.7715
2.3351 | -80.7796
2.3358 | -80.7863
2.3366 | -80.7891
2.3374 | -80.7938
2.3382 | -80.7993
2.3389 | -80.8038
2.3397 | -80.8115
2.3405 | -80.8145
2.3413 | -80.8183
2.342 | -80.8273
2.3428 | -80.8341
2.3436 | -80.8403
2.3444 | -80.8475
2.3451 | -80.8623
2.3459 | -80.8697
2.3467 | -80.8776
2.3475 | -80.8845
2.3482 | -80.8913
2.349 | -80.8948
2.3498 | -80.8989
2.3506 | -80.9048
2.3513 | -80.9112
2.3521 | -80.9193
2.3529 | -80.9279
2.3537 | -80.937
2.3544 | -80.9456
2.3552 | -80.9501
2.356 | -80.9522
2.3568 | -80.9562
2.3575 | -80.9667
2.3583 | -80.9753
2.3591 | -80.983
2.3598 | -80.9918
2.3606 | -81.0013
2.3614 | -81.0076
2.3622 | -81.009
2.3629 | -81.0166
2.3637 | -81.0242
2.3645 | -81.0285
2.3653 | -81.036
2.366 | -81.0452
2.3668 | -81.0599
2.3676 | -81.0755
2.3684 | -81.0872
2.3691 | -81.0981
2.3699 | -81.1081
2.3707 | -81.1168
2.3715 | -81.1258
2.3722 | -81.1357
2.373 | -81.1466
2.3738 | -81.1558
2.3746 | -81.1652
2.3753 | -81.176
2.3761 | -81.1872
2.3769 | -81.1961
2.3777 | -81.2071
2.3784 | -81.2222
2.3792 | -81.236
2.38 | -81.2484
2.3808 | -81.2584
2.3815 | -81.2641
2.3823 | -81.2704
2.3831 | -81.2862
2.3839 | -81.2922
2.3846 | -81.2965
2.3854 | -81.3041
2.3862 | -81.3114
2.387 | -81.3142
2.3877 | -81.3202
2.3885 | -81.3278
2.3893 | -81.3365
2.3901 | -81.3451
2.3908 | -81.3557
2.3916 | -81.3646
2.3924 | -81.379
2.3932 | -81.3894
2.3939 | -81.3956
2.3947 | -81.4051
2.3955 | -81.4153
2.3963 | -81.4195
2.397 | -81.427
2.3978 | -81.437
2.3986 | -81.4489
2.3994 | -81.4562
2.4001 | -81.4582
2.4009 | -81.4638
2.4017 | -81.4697
2.4025 | -81.4748
2.4032 | -81.4821
2.404 | -81.4837
2.4048 | -81.4911
2.4056 | -81.4984
2.4063 | -81.5021
2.4071 | -81.508
2.4079 | -81.5127
2.4087 | -81.52
2.4094 | -81.5269
2.4102 | -81.5303
2.411 | -81.5334
2.4118 | -81.5336
2.4125 | -81.5363
2.4133 | -81.5387
2.4141 | -81.5342
2.4149 | -81.5384
2.4156 | -81.5396
2.4164 | -81.5394
2.4172 | -81.5486
2.418 | -81.5618
2.4187 | -81.5686
2.4195 | -81.5688
2.4203 | -81.5687
2.4211 | -81.5695
2.4218 | -81.5721
2.4226 | -81.5741
2.4234 | -81.5741
2.4242 | -81.5716
2.4249 | -81.5691
2.4257 | -81.5654
2.4265 | -81.5634
2.4273 | -81.5615
2.428 | -81.5636
2.4288 | -81.5651
2.4296 | -81.5612
2.4304 | -81.5597
2.4311 | -81.5623
2.4319 | -81.5614
2.4327 | -81.554
2.4335 | -81.5527
2.4342 | -81.5513
2.435 | -81.547
2.4358 | -81.5423
2.4365 | -81.5389
2.4373 | -81.5336
2.4381 | -81.5299
2.4389 | -81.5269
2.4396 | -81.5174
2.4404 | -81.5093
2.4412 | -81.5029
2.442 | -81.4931
2.4427 | -81.4801
2.4435 | -81.4679
2.4443 | -81.4599
2.4451 | -81.4552
2.4458 | -81.4469
2.4466 | -81.4366
2.4474 | -81.4293
2.4482 | -81.4219
2.4489 | -81.4149
2.4497 | -81.4105
2.4505 | -81.4048
2.4513 | -81.394
2.452 | -81.382
2.4528 | -81.3761
2.4536 | -81.3698
2.4544 | -81.3642
2.4551 | -81.3582
2.4559 | -81.3492
2.4567 | -81.3373
2.4575 | -81.3322
2.4582 | -81.33
2.459 | -81.3283
2.4598 | -81.3286
2.4606 | -81.3251
2.4613 | -81.3193
2.4621 | -81.3139
2.4629 | -81.3117
2.4637 | -81.3101
2.4644 | -81.309
2.4652 | -81.3055
2.466 | -81.3025
2.4668 | -81.2978
2.4675 | -81.2941
2.4683 | -81.2951
2.4691 | -81.2964
2.4699 | -81.292
2.4706 | -81.29
2.4714 | -81.2934
2.4722 | -81.2939
2.473 | -81.29
2.4737 | -81.2834
2.4745 | -81.2806
2.4753 | -81.2795
2.4761 | -81.2783
2.4768 | -81.2726
2.4776 | -81.2719
2.4784 | -81.272
2.4792 | -81.268
2.4799 | -81.266
2.4807 | -81.2692
2.4815 | -81.2684
2.4823 | -81.2647
2.483 | -81.259
2.4838 | -81.2572
2.4846 | -81.2581
2.4854 | -81.2543
2.4861 | -81.249
2.4869 | -81.2394
2.4877 | -81.2319
2.4885 | -81.2287
2.4892 | -81.2294
2.49 | -81.2303
2.4908 | -81.2306
2.4916 | -81.2272
2.4923 | -81.2263
2.4931 | -81.226
2.4939 | -81.226
2.4947 | -81.2242
2.4954 | -81.2226
2.4962 | -81.2226
2.497 | -81.2257
2.4978 | -81.2275
2.4985 | -81.231
2.4993 | -81.2351
2.5001 | -81.2378
2.5009 | -81.2374
2.5016 | -81.2378
2.5024 | -81.2354
2.5032 | -81.2395
2.504 | -81.2426
2.5047 | -81.2449
2.5055 | -81.2466
2.5063 | -81.2509
2.5071 | -81.2516
2.5078 | -81.2527
2.5086 | -81.2548
2.5094 | -81.2577
2.5101 | -81.2614
2.5109 | -81.2598
2.5117 | -81.2585
2.5125 | -81.2594
2.5132 | -81.2582
2.514 | -81.2636
2.5148 | -81.2626
2.5156 | -81.2645
2.5163 | -81.2675
2.5171 | -81.2697
2.5179 | -81.2709
2.5187 | -81.2738
2.5194 | -81.2792
2.5202 | -81.2864
2.521 | -81.29
2.5218 | -81.2939
2.5225 | -81.2968
2.5233 | -81.2981
2.5241 | -81.2994
2.5249 | -81.3025
2.5256 | -81.3039
2.5264 | -81.31
2.5272 | -81.32
2.528 | -81.3283
2.5287 | -81.3319
2.5295 | -81.3324
2.5303 | -81.3299
2.5311 | -81.3276
2.5318 | -81.3259
2.5326 | -81.3277
2.5334 | -81.3369
2.5342 | -81.3363
2.5349 | -81.3396
2.5357 | -81.3453
2.5365 | -81.3501
2.5373 | -81.3486
2.538 | -81.3476
2.5388 | -81.3483
2.5396 | -81.3476
2.5404 | -81.3491
2.5411 | -81.3499
2.5419 | -81.3478
2.5427 | -81.35
2.5435 | -81.3521
2.5442 | -81.3507
2.545 | -81.3528
2.5458 | -81.3507
2.5466 | -81.352
2.5473 | -81.357
2.5481 | -81.3535
2.5489 | -81.3601
2.5497 | -81.3657
2.5504 | -81.3685
2.5512 | -81.3688
2.552 | -81.3697
2.5528 | -81.3705
2.5535 | -81.37
2.5543 | -81.3727
2.5551 | -81.3729
2.5559 | -81.3694
2.5566 | -81.3715
2.5574 | -81.375
2.5582 | -81.3704
2.559 | -81.3655
2.5597 | -81.3686
2.5605 | -81.368
2.5613 | -81.3658
2.5621 | -81.363
2.5628 | -81.361
2.5636 | -81.3603
2.5644 | -81.359
2.5652 | -81.36
2.5659 | -81.3583
2.5667 | -81.3602
2.5675 | -81.3594
2.5683 | -81.3577
2.569 | -81.3555
2.5698 | -81.3542
2.5706 | -81.3549
2.5714 | -81.3524
2.5721 | -81.3517
2.5729 | -81.3534
2.5737 | -81.3567
2.5745 | -81.3639
2.5752 | -81.3626
2.576 | -81.3605
2.5768 | -81.3607
2.5776 | -81.361
2.5783 | -81.3642
2.5791 | -81.3638
2.5799 | -81.3638
2.5806 | -81.3636
2.5814 | -81.3564
2.5822 | -81.346
2.583 | -81.3409
2.5837 | -81.34
2.5845 | -81.3386
2.5853 | -81.3383
2.5861 | -81.34
2.5868 | -81.3395
2.5876 | -81.3378
2.5884 | -81.3364
2.5892 | -81.3355
2.5899 | -81.3351
2.5907 | -81.3349
2.5915 | -81.3368
2.5923 | -81.3375
2.593 | -81.3323
2.5938 | -81.3306
2.5946 | -81.3297
2.5954 | -81.33
2.5961 | -81.3267
2.5969 | -81.3208
2.5977 | -81.3134
2.5985 | -81.3131
2.5992 | -81.3129
2.6 | -81.3136
2.6008 | -81.3143
2.6016 | -81.3102
2.6023 | -81.3055
2.6031 | -81.3069
2.6039 | -81.3009
2.6047 | -81.2945
2.6054 | -81.2942
2.6062 | -81.2947
2.607 | -81.2946
2.6078 | -81.2937
2.6085 | -81.2973
2.6093 | -81.2934
2.6101 | -81.2848
2.6109 | -81.2815
2.6116 | -81.2779
2.6124 | -81.2723
2.6132 | -81.268
2.614 | -81.2578
2.6147 | -81.2535
2.6155 | -81.2491
2.6163 | -81.2473
2.6171 | -81.2476
2.6178 | -81.2462
2.6186 | -81.2422
2.6194 | -81.2421
2.6202 | -81.2448
2.6209 | -81.2441
2.6217 | -81.2419
2.6225 | -81.2379
2.6233 | -81.2355
2.624 | -81.2337
2.6248 | -81.2252
2.6256 | -81.2159
2.6264 | -81.2116
2.6271 | -81.2078
2.6279 | -81.2033
2.6287 | -81.203
2.6295 | -81.2028
2.6302 | -81.1976
2.631 | -81.1869
2.6318 | -81.1724
2.6326 | -81.1627
2.6333 | -81.1563
2.6341 | -81.1485
2.6349 | -81.1371
2.6357 | -81.1282
2.6364 | -81.119
2.6372 | -81.1089
2.638 | -81.0991
2.6388 | -81.0916
2.6395 | -81.0832
2.6403 | -81.0748
2.6411 | -81.0642
2.6419 | -81.0541
2.6426 | -81.0448
2.6434 | -81.0341
2.6442 | -81.0254
2.645 | -81.0174
2.6457 | -81.0115
2.6465 | -81.0014
2.6473 | -80.9902
2.6481 | -80.9813
2.6488 | -80.97
2.6496 | -80.9588
2.6504 | -80.9448
2.6512 | -80.9335
2.6519 | -80.924
2.6527 | -80.9116
2.6535 | -80.8959
2.6543 | -80.8836
2.655 | -80.8717
2.6558 | -80.8571
2.6566 | -80.8444
2.6573 | -80.8278
2.6581 | -80.8153
2.6589 | -80.804
2.6597 | -80.7905
2.6604 | -80.7793
2.6612 | -80.763
2.662 | -80.7478
2.6628 | -80.7382
2.6635 | -80.732
2.6643 | -80.7225
2.6651 | -80.711
2.6659 | -80.6993
2.6666 | -80.6876
2.6674 | -80.6793
2.6682 | -80.6689
2.669 | -80.6557
2.6697 | -80.6373
2.6705 | -80.6213
2.6713 | -80.6061
2.6721 | -80.5879
2.6728 | -80.5728
2.6736 | -80.5594
2.6744 | -80.5436
2.6752 | -80.5302
2.6759 | -80.5167
2.6767 | -80.5025
2.6775 | -80.4857
2.6783 | -80.4708
2.679 | -80.4569
2.6798 | -80.443
2.6806 | -80.4294
2.6814 | -80.4172
2.6821 | -80.4078
2.6829 | -80.3977
2.6837 | -80.3885
2.6845 | -80.3759
2.6852 | -80.3624
2.686 | -80.3494
2.6868 | -80.3373
2.6876 | -80.3271
2.6883 | -80.3152
2.6891 | -80.3047
2.6899 | -80.2947
2.6907 | -80.2859
2.6914 | -80.2743
2.6922 | -80.2643
2.693 | -80.2573
2.6938 | -80.2491
2.6945 | -80.2411
2.6953 | -80.2341
2.6961 | -80.2278
2.6969 | -80.2216
2.6976 | -80.2126
2.6984 | -80.2032
2.6992 | -80.194
2.7 | -80.1863
2.7007 | -80.1788
2.7015 | -80.172
2.7023 | -80.1624
2.7031 | -80.1543
2.7038 | -80.1477
2.7046 | -80.1483
2.7054 | -80.1423
2.7062 | -80.138
2.7069 | -80.1347
2.7077 | -80.1264
2.7085 | -80.113
2.7093 | -80.1077
2.71 | -80.104
2.7108 | -80.0959
2.7116 | -80.0852
2.7124 | -80.0735
2.7131 | -80.0697
2.7139 | -80.0701
2.7147 | -80.067
2.7155 | -80.0665
2.7162 | -80.0663
2.717 | -80.0642
2.7178 | -80.0644
2.7186 | -80.0584
2.7193 | -80.054
2.7201 | -80.0521
2.7209 | -80.0532
2.7217 | -80.0554
2.7224 | -80.0545
2.7232 | -80.0559
2.724 | -80.0533
2.7248 | -80.0525
2.7255 | -80.0553
2.7263 | -80.0615
2.7271 | -80.059
2.7279 | -80.0578
2.7286 | -80.0567
2.7294 | -80.0533
2.7302 | -80.055
2.7309 | -80.0578
2.7317 | -80.0594
2.7325 | -80.0635
2.7333 | -80.0669
2.734 | -80.0679
2.7348 | -80.0702
2.7356 | -80.0724
2.7364 | -80.0766
2.7371 | -80.0808
2.7379 | -80.0851
2.7387 | -80.0892
2.7395 | -80.0927
2.7402 | -80.1012
2.741 | -80.1088
2.7418 | -80.1097
2.7426 | -80.1079
2.7433 | -80.1113
2.7441 | -80.1203
2.7449 | -80.1262
2.7457 | -80.1324
2.7464 | -80.1387
2.7472 | -80.1457
2.748 | -80.152
2.7488 | -80.1618
2.7495 | -80.1676
2.7503 | -80.1733
2.7511 | -80.1795
2.7519 | -80.1861
2.7526 | -80.1918
2.7534 | -80.1961
2.7542 | -80.1987
2.755 | -80.2036
2.7557 | -80.2128
2.7565 | -80.22
2.7573 | -80.2258
2.7581 | -80.2331
2.7588 | -80.2349
2.7596 | -80.2316
2.7604 | -80.2366
2.7612 | -80.2398
2.7619 | -80.2426
2.7627 | -80.2466
2.7635 | -80.2509
2.7643 | -80.2558
2.765 | -80.2631
2.7658 | -80.2699
2.7666 | -80.2752
2.7674 | -80.2818
2.7681 | -80.2916
2.7689 | -80.303
2.7697 | -80.3107
2.7705 | -80.3145
2.7712 | -80.3161
2.772 | -80.3218
2.7728 | -80.3261
2.7736 | -80.3315
2.7743 | -80.3363
2.7751 | -80.3424
2.7759 | -80.3505
2.7767 | -80.3568
2.7774 | -80.3614
2.7782 | -80.3669
2.779 | -80.3724
2.7798 | -80.3773
2.7805 | -80.3821
2.7813 | -80.3883
2.7821 | -80.3933
2.7829 | -80.3964
2.7836 | -80.3955
2.7844 | -80.4008
2.7852 | -80.4083
2.786 | -80.4172
2.7867 | -80.4221
2.7875 | -80.4263
2.7883 | -80.4306
2.7891 | -80.4355
2.7898 | -80.4422
2.7906 | -80.4508
2.7914 | -80.4537
2.7922 | -80.4601
2.7929 | -80.4677
2.7937 | -80.4711
2.7945 | -80.472
2.7953 | -80.4751
2.796 | -80.4781
2.7968 | -80.4788
2.7976 | -80.4798
2.7984 | -80.4811
2.7991 | -80.4868
2.7999 | -80.4919
2.8007 | -80.4965
2.8014 | -80.5026
2.8022 | -80.5061
2.803 | -80.5154
2.8038 | -80.52
2.8045 | -80.5224
2.8053 | -80.5256
2.8061 | -80.5279
2.8069 | -80.5311
2.8076 | -80.5332
2.8084 | -80.5335
2.8092 | -80.5326
2.81 | -80.5321
2.8107 | -80.5337
2.8115 | -80.5359
2.8123 | -80.5385
2.8131 | -80.5403
2.8138 | -80.5426
2.8146 | -80.5482
2.8154 | -80.5548
2.8162 | -80.5488
2.8169 | -80.5459
2.8177 | -80.5464
2.8185 | -80.5464
2.8193 | -80.5458
2.82 | -80.5458
2.8208 | -80.5422
2.8216 | -80.5415
2.8224 | -80.5505
2.8231 | -80.5553
2.8239 | -80.5597
2.8247 | -80.566
2.8255 | -80.5709
2.8262 | -80.5725
2.827 | -80.5747
2.8278 | -80.5752
2.8286 | -80.578
2.8293 | -80.5814
2.8301 | -80.5837
2.8309 | -80.5854
2.8317 | -80.585
2.8324 | -80.5856
2.8332 | -80.5876
2.834 | -80.5875
2.8348 | -80.5901
2.8355 | -80.5899
2.8363 | -80.587
2.8371 | -80.5871
2.8379 | -80.5905
2.8386 | -80.589
2.8394 | -80.5896
2.8402 | -80.5838
2.841 | -80.5857
2.8417 | -80.5855
2.8425 | -80.5909
2.8433 | -80.5949
2.8441 | -80.5944
2.8448 | -80.5989
2.8456 | -80.604
2.8464 | -80.6029
2.8472 | -80.6024
2.8479 | -80.6067
2.8487 | -80.6085
2.8495 | -80.611
2.8503 | -80.6119
2.851 | -80.6104
2.8518 | -80.6081
2.8526 | -80.6066
2.8534 | -80.606
2.8541 | -80.6081
2.8549 | -80.6057
2.8557 | -80.601
2.8565 | -80.6005
2.8572 | -80.6004
2.858 | -80.5991
2.8588 | -80.5981
2.8596 | -80.598
2.8603 | -80.5955
2.8611 | -80.5905
2.8619 | -80.5916
2.8627 | -80.5913
2.8634 | -80.5912
2.8642 | -80.5877
2.865 | -80.5847
2.8658 | -80.5886
2.8665 | -80.5877
2.8673 | -80.5823
2.8681 | -80.5813
2.8689 | -80.5799
2.8696 | -80.5776
2.8704 | -80.5758
2.8712 | -80.5714
2.872 | -80.5656
2.8727 | -80.5585
2.8735 | -80.5529
2.8743 | -80.5497
2.875 | -80.5472
2.8758 | -80.5435
2.8766 | -80.5394
2.8774 | -80.5288
2.8781 | -80.5183
2.8789 | -80.5145
2.8797 | -80.5115
2.8805 | -80.5082
2.8812 | -80.5045
2.882 | -80.4997
2.8828 | -80.4891
2.8836 | -80.4828
2.8843 | -80.4788
2.8851 | -80.4737
2.8859 | -80.4632
2.8867 | -80.4583
2.8874 | -80.4546
2.8882 | -80.4501
2.889 | -80.4456
2.8898 | -80.439
2.8905 | -80.4347
2.8913 | -80.4294
2.8921 | -80.4242
2.8929 | -80.4197
2.8936 | -80.4157
2.8944 | -80.412
2.8952 | -80.4072
2.896 | -80.4018
2.8967 | -80.3968
2.8975 | -80.3912
2.8983 | -80.3812
2.8991 | -80.3745
2.8998 | -80.3699
2.9006 | -80.3664
2.9014 | -80.3601
2.9022 | -80.353
2.9029 | -80.3418
2.9037 | -80.3338
2.9045 | -80.3353
2.9053 | -80.3314
2.906 | -80.3263
2.9068 | -80.3193
2.9076 | -80.3178
2.9084 | -80.319
2.9091 | -80.3182
2.9099 | -80.312
2.9107 | -80.3035
2.9115 | -80.298
2.9122 | -80.2902
2.913 | -80.2832
2.9138 | -80.2748
2.9146 | -80.2673
2.9153 | -80.2649
2.9161 | -80.2611
2.9169 | -80.2557
2.9177 | -80.2503
2.9184 | -80.2419
2.9192 | -80.234
2.92 | -80.227
2.9208 | -80.2244
2.9215 | -80.2225
2.9223 | -80.2173
2.9231 | -80.2128
2.9239 | -80.2111
2.9246 | -80.2078
2.9254 | -80.2033
2.9262 | -80.2031
2.927 | -80.2044
2.9277 | -80.2065
2.9285 | -80.2047
2.9293 | -80.1991
2.9301 | -80.1965
2.9308 | -80.1937
2.9316 | -80.1901
2.9324 | -80.1853
2.9332 | -80.1819
2.9339 | -80.1809
2.9347 | -80.1788
2.9355 | -80.1761
2.9363 | -80.1742
2.937 | -80.1732
2.9378 | -80.1728
2.9386 | -80.1693
2.9394 | -80.1665
2.9401 | -80.1674
2.9409 | -80.1723
2.9417 | -80.1758
2.9425 | -80.1747
2.9432 | -80.173
2.944 | -80.1712
2.9448 | -80.1715
2.9456 | -80.1712
2.9463 | -80.1698
2.9471 | -80.1687
2.9479 | -80.1701
2.9487 | -80.1717
2.9494 | -80.1726
2.9502 | -80.18
2.951 | -80.1821
2.9517 | -80.185
2.9525 | -80.1875
2.9533 | -80.1924
2.9541 | -80.198
2.9548 | -80.2008
2.9556 | -80.2027
2.9564 | -80.205
2.9572 | -80.2098
2.9579 | -80.2145
2.9587 | -80.2234
2.9595 | -80.2359
2.9603 | -80.2406
2.961 | -80.2403
2.9618 | -80.2402
2.9626 | -80.2448
2.9634 | -80.2492
2.9641 | -80.2522
2.9649 | -80.2555
2.9657 | -80.2596
2.9665 | -80.264
2.9672 | -80.2695
2.968 | -80.2752
2.9688 | -80.2805
2.9696 | -80.2852
2.9703 | -80.2921
2.9711 | -80.2979
2.9719 | -80.3033
2.9727 | -80.3137
2.9734 | -80.3197
2.9742 | -80.3261
2.975 | -80.3326
2.9758 | -80.3382
2.9765 | -80.3458
2.9773 | -80.353
2.9781 | -80.3594
2.9789 | -80.3655
2.9796 | -80.372
2.9804 | -80.3762
2.9812 | -80.3773
2.982 | -80.3798
2.9827 | -80.3849
2.9835 | -80.3905
2.9843 | -80.3976
2.9851 | -80.4096
2.9858 | -80.4177
2.9866 | -80.4233
2.9874 | -80.4289
2.9882 | -80.4364
2.9889 | -80.4423
2.9897 | -80.45
2.9905 | -80.4577
2.9913 | -80.4665
2.992 | -80.4737
2.9928 | -80.4817
2.9936 | -80.4888
2.9944 | -80.498
2.9951 | -80.509
2.9959 | -80.518
2.9967 | -80.5264
2.9975 | -80.5372
2.9982 | -80.5455
2.999 | -80.5539
2.9998 | -80.5644
3.0006 | -80.5747
3.0013 | -80.5893
3.0021 | -80.6004
3.0029 | -80.6085
3.0037 | -80.6157
3.0044 | -80.6227
3.0052 | -80.6315
3.006 | -80.639
3.0068 | -80.6469
3.0075 | -80.6554
3.0083 | -80.6636
3.0091 | -80.6694
3.0099 | -80.6776
3.0106 | -80.6887
3.0114 | -80.6996
3.0122 | -80.7103
3.013 | -80.7214
3.0137 | -80.7295
3.0145 | -80.7372
3.0153 | -80.7452
3.0161 | -80.7547
3.0168 | -80.7636
3.0176 | -80.7737
3.0184 | -80.783
3.0192 | -80.792
3.0199 | -80.8008
3.0207 | -80.8156
3.0215 | -80.8251
3.0223 | -80.8288
3.023 | -80.8366
3.0238 | -80.8454
3.0246 | -80.855
3.0253 | -80.8686
3.0261 | -80.883
3.0269 | -80.8901
3.0277 | -80.8957
3.0284 | -80.8988
3.0292 | -80.9033
3.03 | -80.9068
3.0308 | -80.9138
3.0315 | -80.9191
3.0323 | -80.921
3.0331 | -80.9223
3.0339 | -80.9258
3.0346 | -80.931
3.0354 | -80.9365
3.0362 | -80.9439
3.037 | -80.954
3.0377 | -80.9544
3.0385 | -80.9557
3.0393 | -80.9605
3.0401 | -80.9668
3.0408 | -80.9746
3.0416 | -80.9832
3.0424 | -80.9866
3.0432 | -80.9893
3.0439 | -80.9949
3.0447 | -80.9998
3.0455 | -81.0053
3.0463 | -81.0167
3.047 | -81.0217
3.0478 | -81.0274
3.0486 | -81.029
3.0494 | -81.0315
3.0501 | -81.0361
3.0509 | -81.0371
3.0517 | -81.0387
3.0525 | -81.044
3.0532 | -81.0474
3.054 | -81.0506
3.0548 | -81.0509
3.0556 | -81.0543
3.0563 | -81.0578
3.0571 | -81.0611
3.0579 | -81.0627
3.0587 | -81.0648
3.0594 | -81.0635
3.0602 | -81.0675
3.061 | -81.0745
3.0618 | -81.0751
3.0625 | -81.0788
3.0633 | -81.0789
3.0641 | -81.08
3.0649 | -81.0811
3.0656 | -81.0832
3.0664 | -81.0849
3.0672 | -81.0881
3.068 | -81.0927
3.0687 | -81.088
3.0695 | -81.0816
3.0703 | -81.0786
3.0711 | -81.0762
3.0718 | -81.0776
3.0726 | -81.0803
3.0734 | -81.0789
3.0742 | -81.0784
3.0749 | -81.0802
3.0757 | -81.0801
3.0765 | -81.0788
3.0773 | -81.0791
3.078 | -81.079
3.0788 | -81.0823
3.0796 | -81.0858
3.0804 | -81.0892
3.0811 | -81.0907
3.0819 | -81.0876
3.0827 | -81.0822
3.0835 | -81.0787
3.0842 | -81.0763
3.085 | -81.0764
3.0858 | -81.0815
3.0866 | -81.0859
3.0873 | -81.0831
3.0881 | -81.0799
3.0889 | -81.0773
3.0897 | -81.0737
3.0904 | -81.0701
3.0912 | -81.0688
3.092 | -81.0662
3.0928 | -81.0637
3.0935 | -81.06
3.0943 | -81.0521
3.0951 | -81.0513
3.0959 | -81.046
3.0966 | -81.0368
3.0974 | -81.0318
3.0982 | -81.0253
3.0989 | -81.0232
3.0997 | -81.0144
3.1005 | -81.0094
3.1013 | -81.0017
3.102 | -80.9973
3.1028 | -80.9935
3.1036 | -80.9882
3.1044 | -80.9823
3.1051 | -80.9758
3.1059 | -80.969
3.1067 | -80.9667
3.1075 | -80.9651
3.1082 | -80.9593
3.109 | -80.9535
3.1098 | -80.949
3.1106 | -80.9445
3.1113 | -80.9415
3.1121 | -80.9385
3.1129 | -80.9335
3.1137 | -80.9316
3.1144 | -80.9268
3.1152 | -80.9212
3.116 | -80.9106
3.1168 | -80.9025
3.1175 | -80.8958
3.1183 | -80.8903
3.1191 | -80.8842
3.1199 | -80.8806
3.1206 | -80.8779
3.1214 | -80.8753
3.1222 | -80.8713
3.123 | -80.865
3.1237 | -80.8604
3.1245 | -80.8571
3.1253 | -80.8529
3.1261 | -80.8469
3.1268 | -80.841
3.1276 | -80.8348
3.1284 | -80.8294
3.1292 | -80.8247
3.1299 | -80.8195
3.1307 | -80.8183
3.1315 | -80.8085
3.1323 | -80.7958
3.133 | -80.7903
3.1338 | -80.7785
3.1346 | -80.7672
3.1354 | -80.7584
3.1361 | -80.7534
3.1369 | -80.7485
3.1377 | -80.7406
3.1385 | -80.7309
3.1392 | -80.7252
3.14 | -80.7194
3.1408 | -80.7146
3.1416 | -80.7111
3.1423 | -80.7049
3.1431 | -80.6968
3.1439 | -80.6897
3.1447 | -80.6862
3.1454 | -80.6771
3.1462 | -80.6684
3.147 | -80.6627
3.1478 | -80.656
3.1485 | -80.6367
3.1493 | -80.6265
3.1501 | -80.6217
3.1509 | -80.618
3.1516 | -80.6111
3.1524 | -80.6062
3.1532 | -80.6013
3.154 | -80.5926
3.1547 | -80.5821
3.1555 | -80.5717
3.1563 | -80.5652
3.1571 | -80.5599
3.1578 | -80.5534
3.1586 | -80.5472
3.1594 | -80.5428
3.1602 | -80.5382
3.1609 | -80.532
3.1617 | -80.5211
3.1625 | -80.51
3.1633 | -80.5028
3.164 | -80.4957
3.1648 | -80.4882
3.1656 | -80.4807
3.1664 | -80.4728
3.1671 | -80.4616
3.1679 | -80.456
3.1687 | -80.4506
3.1695 | -80.4448
3.1702 | -80.439
3.171 | -80.4294
3.1718 | -80.4148
3.1726 | -80.4036
3.1733 | -80.397
3.1741 | -80.3908
3.1749 | -80.3856
3.1756 | -80.3847
3.1764 | -80.3834
3.1772 | -80.3774
3.178 | -80.371
3.1787 | -80.3656
3.1795 | -80.3587
3.1803 | -80.352
3.1811 | -80.3436
3.1818 | -80.336
3.1826 | -80.3284
3.1834 | -80.3201
3.1842 | -80.3137
3.1849 | -80.3064
3.1857 | -80.3012
3.1865 | -80.2939
3.1873 | -80.2856
3.188 | -80.2756
3.1888 | -80.2644
3.1896 | -80.2569
3.1904 | -80.2486
3.1911 | -80.2417
3.1919 | -80.2328
3.1927 | -80.2245
3.1935 | -80.2178
3.1942 | -80.2118
3.195 | -80.2055
3.1958 | -80.2014
3.1966 | -80.1981
3.1973 | -80.198
3.1981 | -80.1967
3.1989 | -80.1929
3.1997 | -80.1857
3.2004 | -80.1787
3.2012 | -80.1702
3.202 | -80.1618
3.2028 | -80.157
3.2035 | -80.1524
3.2043 | -80.1444
3.2051 | -80.1378
3.2059 | -80.1309
3.2066 | -80.1187
3.2074 | -80.1013
3.2082 | -80.0899
3.209 | -80.0842
3.2097 | -80.0795
3.2105 | -80.0722
3.2113 | -80.0625
3.2121 | -80.0554
3.2128 | -80.0483
3.2136 | -80.0421
3.2144 | -80.0366
3.2152 | -80.0314
3.2159 | -80.0245
3.2167 | -80.0184
3.2175 | -80.0123
3.2183 | -80.0074
3.219 | -79.9999
3.2198 | -79.9942
3.2206 | -79.989
3.2214 | -79.9841
3.2221 | -79.9805
3.2229 | -79.9767
3.2237 | -79.972
3.2245 | -79.9654
3.2252 | -79.961
3.226 | -79.9582
3.2268 | -79.9572
3.2276 | -79.9538
3.2283 | -79.9476
3.2291 | -79.9444
3.2299 | -79.9409
3.2307 | -79.9366
3.2314 | -79.9345
3.2322 | -79.9322
3.233 | -79.9293
3.2338 | -79.9251
3.2345 | -79.9233
3.2353 | -79.9194
3.2361 | -79.9166
3.2369 | -79.9112
3.2376 | -79.9083
3.2384 | -79.9031
3.2392 | -79.8928
3.24 | -79.8928
3.2407 | -79.8874
3.2415 | -79.8889
3.2423 | -79.8875
3.2431 | -79.8856
3.2438 | -79.883
3.2446 | -79.8793
3.2454 | -79.8772
3.2462 | -79.8758
3.2469 | -79.8759
3.2477 | -79.8775
3.2485 | -79.8774
3.2492 | -79.8792
3.25 | -79.8782
3.2508 | -79.8778
3.2516 | -79.8757
3.2523 | -79.8723
3.2531 | -79.8709
3.2539 | -79.8702
3.2547 | -79.8667
3.2554 | -79.8669
3.2562 | -79.8643
3.257 | -79.8671
3.2578 | -79.8702
3.2585 | -79.8691
3.2593 | -79.8674
3.2601 | -79.8657
3.2609 | -79.864
3.2616 | -79.863
3.2624 | -79.8636
3.2632 | -79.862
3.264 | -79.8632
3.2647 | -79.8634
3.2655 | -79.8652
3.2663 | -79.8665
3.2671 | -79.8656
3.2678 | -79.8659
3.2686 | -79.866
3.2694 | -79.8686
3.2702 | -79.8701
3.2709 | -79.871
3.2717 | -79.8742
3.2725 | -79.8758
3.2733 | -79.877
3.274 | -79.8789
3.2748 | -79.8809
3.2756 | -79.885
3.2764 | -79.8881
3.2771 | -79.8884
3.2779 | -79.8889
3.2787 | -79.8925
3.2795 | -79.8947
3.2802 | -79.8979
3.281 | -79.8975
3.2818 | -79.8982
3.2826 | -79.9008
3.2833 | -79.9016
3.2841 | -79.9035
3.2849 | -79.9035
3.2857 | -79.9099
3.2864 | -79.9126
3.2872 | -79.9154
3.288 | -79.9183
3.2888 | -79.9197
3.2895 | -79.9222
3.2903 | -79.9213
3.2911 | -79.9259
3.2919 | -79.9328
3.2926 | -79.9364
3.2934 | -79.939
3.2942 | -79.9426
3.295 | -79.9454
3.2957 | -79.949
3.2965 | -79.9519
3.2973 | -79.9531
3.2981 | -79.9548
3.2988 | -79.9579
3.2996 | -79.9599
3.3004 | -79.9628
3.3012 | -79.9666
3.3019 | -79.9648
3.3027 | -79.9588
3.3035 | -79.9591
3.3043 | -79.9633
3.305 | -79.9685
3.3058 | -79.9726
3.3066 | -79.9761
3.3074 | -79.9788
3.3081 | -79.9831
3.3089 | -79.9865
3.3097 | -79.9881
3.3105 | -79.9922
3.3112 | -79.996
3.312 | -79.9997
3.3128 | -80.0035
3.3136 | -80.0079
3.3143 | -80.013
3.3151 | -80.0181
3.3159 | -80.0249
3.3167 | -80.03
3.3174 | -80.0354
3.3182 | -80.045
3.319 | -80.0505
3.3197 | -80.0535
3.3205 | -80.0576
3.3213 | -80.0633
3.3221 | -80.0682
3.3228 | -80.0694
3.3236 | -80.0773
3.3244 | -80.0808
3.3252 | -80.0872
3.3259 | -80.0897
3.3267 | -80.094
3.3275 | -80.1001
3.3283 | -80.1058
3.329 | -80.1111
3.3298 | -80.1163
3.3306 | -80.1228
3.3314 | -80.1233
3.3321 | -80.1257
3.3329 | -80.1302
3.3337 | -80.14
3.3345 | -80.1433
3.3352 | -80.144
3.336 | -80.1482
3.3368 | -80.1521
3.3376 | -80.1559
3.3383 | -80.1651
3.3391 | -80.1718
3.3399 | -80.178
3.3407 | -80.1829
3.3414 | -80.1921
3.3422 | -80.199
3.343 | -80.204
3.3438 | -80.2076
3.3445 | -80.2125
3.3453 | -80.2182
3.3461 | -80.2229
3.3469 | -80.2276
3.3476 | -80.2317
3.3484 | -80.2341
3.3492 | -80.2377
3.35 | -80.2421
3.3507 | -80.2427
3.3515 | -80.2416
3.3523 | -80.2381
3.3531 | -80.2438
3.3538 | -80.2475
3.3546 | -80.2505
3.3554 | -80.2536
3.3562 | -80.2592
3.3569 | -80.2638
3.3577 | -80.2647
3.3585 | -80.274
3.3593 | -80.2835
3.36 | -80.2843
3.3608 | -80.2929
3.3616 | -80.3006
3.3624 | -80.3051
3.3631 | -80.3101
3.3639 | -80.315
3.3647 | -80.3203
3.3655 | -80.3267
3.3662 | -80.33
3.367 | -80.3338
3.3678 | -80.3373
3.3686 | -80.3406
3.3693 | -80.341
3.3701 | -80.3399
3.3709 | -80.3407
3.3717 | -80.3447
3.3724 | -80.3517
3.3732 | -80.3568
3.374 | -80.3601
3.3748 | -80.3649
3.3755 | -80.3693
3.3763 | -80.3826
3.3771 | -80.3898
3.3779 | -80.3913
3.3786 | -80.3907
3.3794 | -80.3894
3.3802 | -80.3908
3.381 | -80.3933
3.3817 | -80.3949
3.3825 | -80.3993
3.3833 | -80.405
3.3841 | -80.4105
3.3848 | -80.412
3.3856 | -80.409
3.3864 | -80.4044
3.3872 | -80.4029
3.3879 | -80.4007
3.3887 | -80.4018
3.3895 | -80.4039
3.3903 | -80.4077
3.391 | -80.4139
3.3918 | -80.4153
3.3926 | -80.4153
3.3933 | -80.4163
3.3941 | -80.4169
3.3949 | -80.4192
3.3957 | -80.4236
3.3964 | -80.4216
3.3972 | -80.4186
3.398 | -80.4178
3.3988 | -80.4187
3.3995 | -80.4173
3.4003 | -80.4175
3.4011 | -80.4236
3.4019 | -80.4267
3.4026 | -80.4275
3.4034 | -80.4263
3.4042 | -80.426
3.405 | -80.4253
3.4057 | -80.4249
3.4065 | -80.4258
3.4073 | -80.4252
3.4081 | -80.4251
3.4088 | -80.4235
3.4096 | -80.4171
3.4104 | -80.4093
3.4112 | -80.4089
3.4119 | -80.4097
3.4127 | -80.4094
3.4135 | -80.4086
3.4143 | -80.4004
3.415 | -80.3931
3.4158 | -80.39
3.4166 | -80.3889
3.4174 | -80.3866
3.4181 | -80.3876
3.4189 | -80.3895
3.4197 | -80.3871
3.4205 | -80.3857
3.4212 | -80.3844
3.422 | -80.3815
3.4228 | -80.3816
3.4236 | -80.3803
3.4243 | -80.3779
3.4251 | -80.3758
3.4259 | -80.3745
3.4267 | -80.3732
3.4274 | -80.3703
3.4282 | -80.3671
3.429 | -80.3638
3.4298 | -80.3623
3.4305 | -80.3606
3.4313 | -80.3616
3.4321 | -80.3587
3.4329 | -80.3551
3.4336 | -80.3534
3.4344 | -80.3492
3.4352 | -80.3422
3.436 | -80.3394
3.4367 | -80.3371
3.4375 | -80.3343
3.4383 | -80.3376
3.4391 | -80.3462
3.4398 | -80.3476
3.4406 | -80.3432
3.4414 | -80.34
3.4422 | -80.3386
3.4429 | -80.3305
3.4437 | -80.3241
3.4445 | -80.32
3.4453 | -80.3169
3.446 | -80.3118
3.4468 | -80.306
3.4476 | -80.2997
3.4484 | -80.2951
3.4491 | -80.2911
3.4499 | -80.2864
3.4507 | -80.283
3.4515 | -80.2824
3.4522 | -80.2803
3.453 | -80.276
3.4538 | -80.2713
3.4546 | -80.267
3.4553 | -80.2577
3.4561 | -80.2492
3.4569 | -80.2465
3.4577 | -80.2459
3.4584 | -80.2408
3.4592 | -80.2352
3.46 | -80.231
3.4608 | -80.2286
3.4615 | -80.2259
3.4623 | -80.2228
3.4631 | -80.2196
3.4639 | -80.2167
3.4646 | -80.2135
3.4654 | -80.2098
3.4662 | -80.2028
3.467 | -80.1939
3.4677 | -80.1865
3.4685 | -80.1809
3.4693 | -80.179
3.47 | -80.1759
3.4708 | -80.1746
3.4716 | -80.1694
3.4724 | -80.1699
3.4731 | -80.1752
3.4739 | -80.1725
3.4747 | -80.1704
3.4755 | -80.1706
3.4762 | -80.1671
3.477 | -80.162
3.4778 | -80.1583
3.4786 | -80.1549
3.4793 | -80.1519
3.4801 | -80.1435
3.4809 | -80.1365
3.4817 | -80.133
3.4824 | -80.1294
3.4832 | -80.1264
3.484 | -80.1232
3.4848 | -80.1202
3.4855 | -80.1172
3.4863 | -80.1157
3.4871 | -80.1188
3.4879 | -80.1206
3.4886 | -80.1191
3.4894 | -80.1164
3.4902 | -80.1162
3.491 | -80.1133
3.4917 | -80.1125
3.4925 | -80.1114
3.4933 | -80.1101
3.4941 | -80.1066
3.4948 | -80.1028
3.4956 | -80.1013
3.4964 | -80.1002
3.4972 | -80.0987
3.4979 | -80.0972
3.4987 | -80.0982
3.4995 | -80.0981
3.5003 | -80.0963
3.501 | -80.0957
3.5018 | -80.092
3.5026 | -80.0874
3.5034 | -80.0869
3.5041 | -80.0859
3.5049 | -80.0863
3.5057 | -80.0843
3.5065 | -80.0803
3.5072 | -80.08
3.508 | -80.0767
3.5088 | -80.0798
3.5096 | -80.0807
3.5103 | -80.0813
3.5111 | -80.0786
3.5119 | -80.077
3.5127 | -80.0771
3.5134 | -80.0772
3.5142 | -80.0782
3.515 | -80.078
3.5158 | -80.0778
3.5165 | -80.0779
3.5173 | -80.0798
3.5181 | -80.0821
3.5189 | -80.0826
3.5196 | -80.0822
3.5204 | -80.0819
3.5212 | -80.0796
3.522 | -80.0804
3.5227 | -80.0788
3.5235 | -80.0788
3.5243 | -80.0798
3.5251 | -80.0807
3.5258 | -80.0795
3.5266 | -80.0793
3.5274 | -80.0812
3.5282 | -80.0867
3.5289 | -80.09
3.5297 | -80.0899
3.5305 | -80.0895
3.5313 | -80.0905
3.532 | -80.0955
3.5328 | -80.093
3.5336 | -80.0895
3.5344 | -80.0854
3.5351 | -80.0773
3.5359 | -80.0733
3.5367 | -80.0696
3.5375 | -80.0694
3.5382 | -80.0713
3.539 | -80.0712
3.5398 | -80.0688
3.5406 | -80.072
3.5413 | -80.0807
3.5421 | -80.0838
3.5429 | -80.0847
3.5436 | -80.0808
3.5444 | -80.0783
3.5452 | -80.077
3.546 | -80.0776
3.5467 | -80.0739
3.5475 | -80.0641
3.5483 | -80.0556
3.5491 | -80.0539
3.5498 | -80.0518
3.5506 | -80.0522
3.5514 | -80.052
3.5522 | -80.0514
3.5529 | -80.0514
3.5537 | -80.0477
3.5545 | -80.044
3.5553 | -80.0417
3.556 | -80.0411
3.5568 | -80.0413
3.5576 | -80.039
3.5584 | -80.0362
3.5591 | -80.0362
3.5599 | -80.037
3.5607 | -80.0386
3.5615 | -80.0379
3.5622 | -80.0382
3.563 | -80.0371
3.5638 | -80.0329
3.5646 | -80.0326
3.5653 | -80.0317
3.5661 | -80.0314
3.5669 | -80.0293
3.5677 | -80.0274
3.5684 | -80.0254
3.5692 | -80.0243
3.57 | -80.0205
3.5708 | -80.0175
3.5715 | -80.0153
3.5723 | -80.0165
3.5731 | -80.0162
3.5739 | -80.0138
3.5746 | -80.0114
3.5754 | -80.0117
3.5762 | -80.008
3.577 | -80.0081
3.5777 | -80.0055
3.5785 | -80.0016
3.5793 | -79.9998
3.5801 | -79.9977
3.5808 | -79.9952
3.5816 | -79.9923
3.5824 | -79.9895
3.5832 | -79.9873
3.5839 | -79.9859
3.5847 | -79.983
3.5855 | -79.9777
3.5863 | -79.9736
3.587 | -79.9657
3.5878 | -79.9591
3.5886 | -79.9541
3.5894 | -79.9494
3.5901 | -79.9488
3.5909 | -79.9448
3.5917 | -79.9405
3.5925 | -79.9404
3.5932 | -79.9422
3.594 | -79.9406
3.5948 | -79.9398
3.5956 | -79.935
3.5963 | -79.9315
3.5971 | -79.928
3.5979 | -79.9246
3.5987 | -79.9125
3.5994 | -79.9062
3.6002 | -79.9029
3.601 | -79.9007
3.6018 | -79.8953
3.6025 | -79.8874
3.6033 | -79.8803
3.6041 | -79.8744
3.6049 | -79.87
3.6056 | -79.865
3.6064 | -79.8579
3.6072 | -79.8524
3.608 | -79.8465
3.6087 | -79.8397
3.6095 | -79.8347
3.6103 | -79.8283
3.6111 | -79.8228
3.6118 | -79.8162
3.6126 | -79.8127
3.6134 | -79.8131
3.6141 | -79.8123
3.6149 | -79.8073
3.6157 | -79.8016
3.6165 | -79.7953
3.6172 | -79.7892
3.618 | -79.7838
3.6188 | -79.7777
3.6196 | -79.773
3.6203 | -79.769
3.6211 | -79.7665
3.6219 | -79.7612
3.6227 | -79.7577
3.6234 | -79.7554
3.6242 | -79.7527
3.625 | -79.7492
3.6258 | -79.746
3.6265 | -79.7421
3.6273 | -79.736
3.6281 | -79.7307
3.6289 | -79.7267
3.6296 | -79.7243
3.6304 | -79.7238
3.6312 | -79.7229
3.632 | -79.7203
3.6327 | -79.7174
3.6335 | -79.7129
3.6343 | -79.7089
3.6351 | -79.7067
3.6358 | -79.7035
3.6366 | -79.7008
3.6374 | -79.6996
3.6382 | -79.6975
3.6389 | -79.6949
3.6397 | -79.6919
3.6405 | -79.6906
3.6413 | -79.6905
3.642 | -79.6903
3.6428 | -79.6944
3.6436 | -79.6941
3.6444 | -79.695
3.6451 | -79.6938
3.6459 | -79.6919
3.6467 | -79.6894
3.6475 | -79.6854
3.6482 | -79.6791
3.649 | -79.6756
3.6498 | -79.6761
3.6506 | -79.674
3.6513 | -79.6733
3.6521 | -79.6732
3.6529 | -79.6717
3.6537 | -79.6682
3.6544 | -79.6655
3.6552 | -79.6648
3.656 | -79.6675
3.6568 | -79.673
3.6575 | -79.6756
3.6583 | -79.6761
3.6591 | -79.6777
3.6599 | -79.6804
3.6606 | -79.6877
3.6614 | -79.6942
3.6622 | -79.697
3.663 | -79.7001
3.6637 | -79.7036
3.6645 | -79.7045
3.6653 | -79.713
3.6661 | -79.7113
3.6668 | -79.712
3.6676 | -79.7145
3.6684 | -79.7179
3.6692 | -79.7215
3.6699 | -79.7258
3.6707 | -79.7296
3.6715 | -79.7343
3.6723 | -79.7393
3.673 | -79.744
3.6738 | -79.7481
3.6746 | -79.7506
3.6754 | -79.7555
3.6761 | -79.7615
3.6769 | -79.7682
3.6777 | -79.7739
3.6785 | -79.78
3.6792 | -79.7862
3.68 | -79.7924
3.6808 | -79.7991
3.6816 | -79.8066
3.6823 | -79.8132
3.6831 | -79.8214
3.6839 | -79.8293
3.6847 | -79.8368
3.6854 | -79.8439
3.6862 | -79.8511
3.687 | -79.8587
3.6877 | -79.8724
3.6885 | -79.8789
3.6893 | -79.8851
3.6901 | -79.8858
3.6908 | -79.8917
3.6916 | -79.8999
3.6924 | -79.9126
3.6932 | -79.9244
3.6939 | -79.9311
3.6947 | -79.9389
3.6955 | -79.9486
3.6963 | -79.9601
3.697 | -79.971
3.6978 | -79.9817
3.6986 | -79.9912
3.6994 | -80.0007
3.7001 | -80.0066
3.7009 | -80.0109
3.7017 | -80.0205
3.7025 | -80.031
3.7032 | -80.0406
3.704 | -80.0445
3.7048 | -80.0542
3.7056 | -80.0668
3.7063 | -80.0737
3.7071 | -80.0809
3.7079 | -80.0875
3.7087 | -80.0979
3.7094 | -80.1071
3.7102 | -80.1174
3.711 | -80.1273
3.7118 | -80.1379
3.7125 | -80.1463
3.7133 | -80.1518
3.7141 | -80.1562
3.7149 | -80.1642
3.7156 | -80.1746
3.7164 | -80.1848
3.7172 | -80.1969
3.718 | -80.209
3.7187 | -80.2203
3.7195 | -80.2298
3.7203 | -80.2376
3.7211 | -80.2449
3.7218 | -80.2532
3.7226 | -80.2629
3.7234 | -80.2715
3.7242 | -80.278
3.7249 | -80.2862
3.7257 | -80.2927
3.7265 | -80.299
3.7273 | -80.307
3.728 | -80.3127
3.7288 | -80.3169
3.7296 | -80.3239
3.7304 | -80.3299
3.7311 | -80.3418
3.7319 | -80.3522
3.7327 | -80.3579
3.7335 | -80.3649
3.7342 | -80.3713
3.735 | -80.378
3.7358 | -80.3833
3.7366 | -80.3875
3.7373 | -80.3937
3.7381 | -80.4002
3.7389 | -80.4024
3.7397 | -80.4045
3.7404 | -80.4099
3.7412 | -80.4151
3.742 | -80.4183
3.7428 | -80.4246
3.7435 | -80.43
3.7443 | -80.4328
3.7451 | -80.4365
3.7459 | -80.4412
3.7466 | -80.4443
3.7474 | -80.4435
3.7482 | -80.4417
3.749 | -80.4422
3.7497 | -80.4452
3.7505 | -80.4456
3.7513 | -80.4458
3.7521 | -80.4482
3.7528 | -80.4509
3.7536 | -80.4534
3.7544 | -80.4558
3.7552 | -80.4536
3.7559 | -80.4589
3.7567 | -80.4609
3.7575 | -80.463
3.7583 | -80.4644
3.759 | -80.4671
3.7598 | -80.4687
3.7606 | -80.4632
3.7614 | -80.4638
3.7621 | -80.4648
3.7629 | -80.4638
3.7637 | -80.4607
3.7644 | -80.4582
3.7652 | -80.46
3.766 | -80.4596
3.7668 | -80.4543
3.7675 | -80.445
3.7683 | -80.4406
3.7691 | -80.4418
3.7699 | -80.4417
3.7706 | -80.4401
3.7714 | -80.439
3.7722 | -80.4365
3.773 | -80.4358
3.7737 | -80.4347
3.7745 | -80.4355
3.7753 | -80.4374
3.7761 | -80.4369
3.7768 | -80.4355
3.7776 | -80.4295
3.7784 | -80.4246
3.7792 | -80.4193
3.7799 | -80.4157
3.7807 | -80.4121
3.7815 | -80.4093
3.7823 | -80.4057
3.783 | -80.4017
3.7838 | -80.3979
3.7846 | -80.3961
3.7854 | -80.3935
3.7861 | -80.3839
3.7869 | -80.3747
3.7877 | -80.3691
3.7885 | -80.3684
3.7892 | -80.3669
3.79 | -80.366
3.7908 | -80.3666
3.7916 | -80.363
3.7923 | -80.3609
3.7931 | -80.3559
3.7939 | -80.3487
3.7947 | -80.3396
3.7954 | -80.3358
3.7962 | -80.3314
3.797 | -80.3256
3.7978 | -80.3188
3.7985 | -80.3122
3.7993 | -80.3063
3.8001 | -80.3043
3.8009 | -80.3
3.8016 | -80.2937
3.8024 | -80.2866
3.8032 | -80.2846
3.804 | -80.2872
3.8047 | -80.2863
3.8055 | -80.2813
3.8063 | -80.2743
3.8071 | -80.2649
3.8078 | -80.2555
3.8086 | -80.2467
3.8094 | -80.2382
3.8102 | -80.2298
3.8109 | -80.2232
3.8117 | -80.2181
3.8125 | -80.2112
3.8133 | -80.2035
3.814 | -80.1958
3.8148 | -80.1889
3.8156 | -80.1835
3.8164 | -80.1745
3.8171 | -80.1686
3.8179 | -80.1603
3.8187 | -80.1517
3.8195 | -80.1448
3.8202 | -80.1402
3.821 | -80.1361
3.8218 | -80.1317
3.8226 | -80.1251
3.8233 | -80.1186
3.8241 | -80.1124
3.8249 | -80.1055
3.8257 | -80.0989
3.8264 | -80.0934
3.8272 | -80.0884
3.828 | -80.086
3.8288 | -80.0822
3.8295 | -80.0767
3.8303 | -80.0717
3.8311 | -80.064
3.8319 | -80.0567
3.8326 | -80.0525
3.8334 | -80.0476
3.8342 | -80.0427
3.835 | -80.0372
3.8357 | -80.0329
3.8365 | -80.0277
3.8373 | -80.0237
3.838 | -80.0194
3.8388 | -80.0151
3.8396 | -80.0114
3.8404 | -80.0081
3.8411 | -80.0048
3.8419 | -80.0009
3.8427 | -79.9962
3.8435 | -79.9923
3.8442 | -79.9894
3.845 | -79.989
3.8458 | -79.9876
3.8466 | -79.9884
3.8473 | -79.9888
3.8481 | -79.9861
3.8489 | -79.9859
3.8497 | -79.9844
3.8504 | -79.9819
3.8512 | -79.9811
3.852 | -79.9796
3.8528 | -79.9781
3.8535 | -79.9784
3.8543 | -79.9771
3.8551 | -79.9754
3.8559 | -79.9714
3.8566 | -79.9707
3.8574 | -79.9684
3.8582 | -79.9699
3.859 | -79.9721
3.8597 | -79.9675
3.8605 | -79.9631
3.8613 | -79.9625
3.8621 | -79.9626
3.8628 | -79.9645
3.8636 | -79.9667
3.8644 | -79.9671
3.8652 | -79.9664
3.8659 | -79.9674
3.8667 | -79.974
3.8675 | -79.9751
3.8683 | -79.9755
3.869 | -79.9777
3.8698 | -79.9805
3.8706 | -79.982
3.8714 | -79.9838
3.8721 | -79.9841
3.8729 | -79.9866
3.8737 | -79.989
3.8745 | -79.9908
3.8752 | -79.9928
3.876 | -79.9944
3.8768 | -79.9978
3.8776 | -80.0021
3.8783 | -80.0051
3.8791 | -80.0049
3.8799 | -80.0058
3.8807 | -80.0066
3.8814 | -80.0085
3.8822 | -80.0098
3.883 | -80.0123
3.8838 | -80.0163
3.8845 | -80.0203
3.8853 | -80.025
3.8861 | -80.0313
3.8869 | -80.0377
3.8876 | -80.0424
3.8884 | -80.0476
3.8892 | -80.056
3.89 | -80.0585
3.8907 | -80.0605
3.8915 | -80.0641
3.8923 | -80.0688
3.8931 | -80.0746
3.8938 | -80.0817
3.8946 | -80.0871
3.8954 | -80.0943
3.8962 | -80.1012
3.8969 | -80.1073
3.8977 | -80.1136
3.8985 | -80.1196
3.8993 | -80.1263
3.9 | -80.1322
3.9008 | -80.1377
3.9016 | -80.1456
3.9024 | -80.1533
3.9031 | -80.1601
3.9039 | -80.1658
3.9047 | -80.1728
3.9055 | -80.1795
3.9062 | -80.1865
3.907 | -80.1944
3.9078 | -80.2015
3.9085 | -80.2063
3.9093 | -80.2136
3.9101 | -80.2243
3.9109 | -80.2362
3.9116 | -80.2453
3.9124 | -80.2579
3.9132 | -80.2673
3.914 | -80.275
3.9147 | -80.2844
3.9155 | -80.2923
3.9163 | -80.3007
3.9171 | -80.3102
3.9178 | -80.3196
3.9186 | -80.3304
3.9194 | -80.338
3.9202 | -80.3452
3.9209 | -80.3531
3.9217 | -80.3566
3.9225 | -80.3629
3.9233 | -80.37
3.924 | -80.3775
3.9248 | -80.3849
3.9256 | -80.3922
3.9264 | -80.4001
3.9271 | -80.4089
3.9279 | -80.4176
3.9287 | -80.4273
3.9295 | -80.4351
3.9302 | -80.4392
3.931 | -80.451
3.9318 | -80.4605
3.9326 | -80.4682
3.9333 | -80.4742
3.9341 | -80.4807
3.9349 | -80.4879
3.9357 | -80.4969
3.9364 | -80.5049
3.9372 | -80.5127
3.938 | -80.5209
3.9388 | -80.5297
3.9395 | -80.5381
3.9403 | -80.5466
3.9411 | -80.5547
3.9419 | -80.5626
3.9426 | -80.5702
3.9434 | -80.5781
3.9442 | -80.5864
3.945 | -80.5945
3.9457 | -80.6049
3.9465 | -80.6135
3.9473 | -80.6164
3.9481 | -80.6175
3.9488 | -80.6252
3.9496 | -80.6327
3.9504 | -80.6405
3.9512 | -80.6465
3.9519 | -80.6469
3.9527 | -80.6504
3.9535 | -80.6565
3.9543 | -80.662
3.955 | -80.6709
3.9558 | -80.6766
3.9566 | -80.6859
3.9574 | -80.6982
3.9581 | -80.7072
3.9589 | -80.7134
3.9597 | -80.7197
3.9605 | -80.7252
3.9612 | -80.7283
3.962 | -80.7308
3.9628 | -80.7356
3.9636 | -80.7397
3.9643 | -80.7436
3.9651 | -80.7487
3.9659 | -80.7536
3.9667 | -80.757
3.9674 | -80.764
3.9682 | -80.7727
3.969 | -80.7765
3.9698 | -80.78
3.9705 | -80.7836
3.9713 | -80.7941
3.9721 | -80.802
3.9729 | -80.8058
3.9736 | -80.8096
3.9744 | -80.8131
3.9752 | -80.8208
3.976 | -80.8236
3.9767 | -80.8269
3.9775 | -80.8299
3.9783 | -80.831
3.9791 | -80.8338
3.9798 | -80.836
3.9806 | -80.8384
3.9814 | -80.8423
3.9821 | -80.8489
3.9829 | -80.8559
3.9837 | -80.8581
3.9845 | -80.8604
3.9852 | -80.8616
3.986 | -80.8623
3.9868 | -80.8661
3.9876 | -80.8706
3.9883 | -80.8734
3.9891 | -80.8755
3.9899 | -80.8781
3.9907 | -80.8792
3.9914 | -80.8796
3.9922 | -80.8737
3.993 | -80.8725
3.9938 | -80.8729
3.9945 | -80.8754
3.9953 | -80.8767
3.9961 | -80.8729
3.9969 | -80.8767
3.9976 | -80.8782
3.9984 | -80.8792
3.9992 | -80.8822
4 | -80.8828
4.0007 | -80.8823
4.0015 | -80.8825
4.0023 | -80.8811
4.0031 | -80.8803
4.0038 | -80.8798
4.0046 | -80.8802
4.0054 | -80.8803
4.0062 | -80.8812
4.0069 | -80.8841
4.0077 | -80.884
4.0085 | -80.8833
4.0093 | -80.8848
4.01 | -80.8856
4.0108 | -80.884
4.0116 | -80.8836
4.0124 | -80.8844
4.0131 | -80.8835
4.0139 | -80.8822
4.0147 | -80.8821
4.0155 | -80.8819
4.0162 | -80.8807
4.017 | -80.8825
4.0178 | -80.8888
4.0186 | -80.8931
4.0193 | -80.8952
4.0201 | -80.8943
4.0209 | -80.8922
4.0217 | -80.8908
4.0224 | -80.8904
4.0232 | -80.8833
4.024 | -80.8809
4.0248 | -80.8786
4.0255 | -80.8789
4.0263 | -80.8782
4.0271 | -80.8742
4.0279 | -80.8683
4.0286 | -80.8664
4.0294 | -80.8641
4.0302 | -80.8612
4.031 | -80.8587
4.0317 | -80.8589
4.0325 | -80.862
4.0333 | -80.8624
4.0341 | -80.8606
4.0348 | -80.8573
4.0356 | -80.858
4.0364 | -80.8561
4.0372 | -80.8481
4.0379 | -80.8441
4.0387 | -80.8464
4.0395 | -80.8516
4.0403 | -80.851
4.041 | -80.8472
4.0418 | -80.8433
4.0426 | -80.841
4.0434 | -80.8382
4.0441 | -80.8356
4.0449 | -80.8331
4.0457 | -80.8318
4.0465 | -80.8291
4.0472 | -80.8247
4.048 | -80.8225
4.0488 | -80.8215
4.0496 | -80.8199
4.0503 | -80.8161
4.0511 | -80.8129
4.0519 | -80.8131
4.0527 | -80.8123
4.0534 | -80.8104
4.0542 | -80.811
4.055 | -80.8112
4.0557 | -80.8098
4.0565 | -80.8046
4.0573 | -80.7994
4.0581 | -80.7967
4.0588 | -80.7962
4.0596 | -80.8027
4.0604 | -80.8045
4.0612 | -80.7992
4.0619 | -80.7922
4.0627 | -80.7897
4.0635 | -80.7886
4.0643 | -80.7892
4.065 | -80.7921
4.0658 | -80.7905
4.0666 | -80.7857
4.0674 | -80.7836
4.0681 | -80.7842
4.0689 | -80.7859
4.0697 | -80.7842
4.0705 | -80.7813
4.0712 | -80.7779
4.072 | -80.7739
4.0728 | -80.7706
4.0736 | -80.7684
4.0743 | -80.7671
4.0751 | -80.7649
4.0759 | -80.7623
4.0767 | -80.7614
4.0774 | -80.761
4.0782 | -80.7602
4.079 | -80.7584
4.0798 | -80.7579
4.0805 | -80.7574
4.0813 | -80.7561
4.0821 | -80.756
4.0829 | -80.7562
4.0836 | -80.7561
4.0844 | -80.7556
4.0852 | -80.7551
4.086 | -80.7559
4.0867 | -80.7559
4.0875 | -80.7548
4.0883 | -80.7516
4.0891 | -80.7499
4.0898 | -80.7503
4.0906 | -80.7501
4.0914 | -80.7477
4.0922 | -80.7489
4.0929 | -80.7478
4.0937 | -80.7484
4.0945 | -80.7494
4.0953 | -80.7485
4.096 | -80.7482
4.0968 | -80.7477
4.0976 | -80.7491
4.0984 | -80.7498
4.0991 | -80.7502
4.0999 | -80.7498
4.1007 | -80.7498
4.1015 | -80.751
4.1022 | -80.7516
4.103 | -80.7495
4.1038 | -80.7446
4.1046 | -80.7425
4.1053 | -80.7427
4.1061 | -80.7434
4.1069 | -80.7434
4.1077 | -80.7465
4.1084 | -80.7492
4.1092 | -80.747
4.11 | -80.7466
4.1108 | -80.7447
4.1115 | -80.7421
4.1123 | -80.7415
4.1131 | -80.7407
4.1139 | -80.7406
4.1146 | -80.7412
4.1154 | -80.7399
4.1162 | -80.7378
4.117 | -80.7374
4.1177 | -80.7373
4.1185 | -80.7366
4.1193 | -80.7375
4.1201 | -80.7396
4.1208 | -80.7425
4.1216 | -80.7484
4.1224 | -80.75
4.1232 | -80.7507
4.1239 | -80.7512
4.1247 | -80.7517
4.1255 | -80.7521
4.1263 | -80.7549
4.127 | -80.7565
4.1278 | -80.7577
4.1286 | -80.7589
4.1293 | -80.7566
4.1301 | -80.7566
4.1309 | -80.7569
4.1317 | -80.7575
4.1324 | -80.7591
4.1332 | -80.7608
4.134 | -80.7644
4.1348 | -80.7676
4.1355 | -80.7684
4.1363 | -80.7713
4.1371 | -80.7721
4.1379 | -80.7751
4.1386 | -80.7815
4.1394 | -80.7846
4.1402 | -80.7853
4.141 | -80.7849
4.1417 | -80.7856
4.1425 | -80.7891
4.1433 | -80.7912
4.1441 | -80.7913
4.1448 | -80.7907
4.1456 | -80.7918
4.1464 | -80.7915
4.1472 | -80.7933
4.1479 | -80.7957
4.1487 | -80.7965
4.1495 | -80.7975
4.1503 | -80.7991
4.151 | -80.8008
4.1518 | -80.8038
4.1526 | -80.807
4.1534 | -80.8095
4.1541 | -80.8103
4.1549 | -80.8107
4.1557 | -80.8122
4.1565 | -80.8138
4.1572 | -80.8181
4.158 | -80.8196
4.1588 | -80.8223
4.1596 | -80.825
4.1603 | -80.8278
4.1611 | -80.8323
4.1619 | -80.8334
4.1627 | -80.8355
4.1634 | -80.8392
4.1642 | -80.8412
4.165 | -80.8414
4.1658 | -80.8416
4.1665 | -80.8427
4.1673 | -80.8439
4.1681 | -80.8456
4.1689 | -80.8474
4.1696 | -80.85
4.1704 | -80.8538
4.1712 | -80.8599
4.172 | -80.8644
4.1727 | -80.8682
4.1735 | -80.8688
4.1743 | -80.8698
4.1751 | -80.8719
4.1758 | -80.8756
4.1766 | -80.8782
4.1774 | -80.8802
4.1782 | -80.8817
4.1789 | -80.8803
4.1797 | -80.8798
4.1805 | -80.8821
4.1813 | -80.8837
4.182 | -80.8846
4.1828 | -80.8874
4.1836 | -80.8896
4.1844 | -80.8924
4.1851 | -80.8951
4.1859 | -80.8986
4.1867 | -80.901
4.1875 | -80.9044
4.1882 | -80.9084
4.189 | -80.9114
4.1898 | -80.9147
4.1906 | -80.9207
4.1913 | -80.9263
4.1921 | -80.9286
4.1929 | -80.9314
4.1937 | -80.9336
4.1944 | -80.937
4.1952 | -80.9415
4.196 | -80.9445
4.1968 | -80.9455
4.1975 | -80.9488
4.1983 | -80.9522
4.1991 | -80.9555
4.1999 | -80.9572
4.2006 | -80.9617
4.2014 | -80.9633
4.2022 | -80.9662
4.203 | -80.9705
4.2037 | -80.9748
4.2045 | -80.976
4.2053 | -80.9803
4.206 | -80.979
4.2068 | -80.9787
4.2076 | -80.9783
4.2084 | -80.9767
4.2091 | -80.9783
4.2099 | -80.981
4.2107 | -80.9821
4.2115 | -80.9846
4.2122 | -80.9867
4.213 | -80.99
4.2138 | -80.9952
4.2146 | -80.9985
4.2153 | -81.0013
4.2161 | -81.0031
4.2169 | -81.0078
4.2177 | -81.0127
4.2184 | -81.0118
4.2192 | -81.0146
4.22 | -81.0165
4.2208 | -81.0188
4.2215 | -81.0208
4.2223 | -81.0226
4.2231 | -81.0237
4.2239 | -81.026
4.2246 | -81.0283
4.2254 | -81.033
4.2262 | -81.0308
4.227 | -81.0312
4.2277 | -81.034
4.2285 | -81.0341
4.2293 | -81.0343
4.2301 | -81.0368
4.2308 | -81.0385
4.2316 | -81.04
4.2324 | -81.0469
4.2332 | -81.0558
4.2339 | -81.0596
4.2347 | -81.0612
4.2355 | -81.0618
4.2363 | -81.0635
4.237 | -81.0713
4.2378 | -81.0765
4.2386 | -81.0793
4.2394 | -81.0813
4.2401 | -81.0811
4.2409 | -81.0825
4.2417 | -81.0847
4.2425 | -81.0866
4.2432 | -81.0897
4.244 | -81.0913
4.2448 | -81.0933
4.2456 | -81.095
4.2463 | -81.0985
4.2471 | -81.1023
4.2479 | -81.1099
4.2487 | -81.114
4.2494 | -81.1147
4.2502 | -81.1175
4.251 | -81.1187
4.2518 | -81.1184
4.2525 | -81.1204
4.2533 | -81.1248
4.2541 | -81.1271
4.2549 | -81.1245
4.2556 | -81.1233
4.2564 | -81.123
4.2572 | -81.1259
4.258 | -81.1284
4.2587 | -81.129
4.2595 | -81.1289
4.2603 | -81.1297
4.2611 | -81.1293
4.2618 | -81.1308
4.2626 | -81.1321
4.2634 | -81.1327
4.2642 | -81.1302
4.2649 | -81.1297
4.2657 | -81.1313
4.2665 | -81.1322
4.2673 | -81.1325
4.268 | -81.1331
4.2688 | -81.1342
4.2696 | -81.1322
4.2704 | -81.1331
4.2711 | -81.1327
4.2719 | -81.1332
4.2727 | -81.1332
4.2735 | -81.1343
4.2742 | -81.1365
4.275 | -81.1371
4.2758 | -81.1354
4.2766 | -81.1356
4.2773 | -81.1359
4.2781 | -81.1346
4.2789 | -81.1366
4.2796 | -81.1386
4.2804 | -81.147
4.2812 | -81.1521
4.282 | -81.1529
4.2827 | -81.1545
4.2835 | -81.1549
4.2843 | -81.1572
4.2851 | -81.1566
4.2858 | -81.1591
4.2866 | -81.1593
4.2874 | -81.1579
4.2882 | -81.1571
4.2889 | -81.156
4.2897 | -81.1552
4.2905 | -81.1552
4.2913 | -81.1559
4.292 | -81.1561
4.2928 | -81.1532
4.2936 | -81.1516
4.2944 | -81.1502
4.2951 | -81.1485
4.2959 | -81.146
4.2967 | -81.1429
4.2975 | -81.1413
4.2982 | -81.1404
4.299 | -81.1403
4.2998 | -81.1413
4.3006 | -81.1413
4.3013 | -81.1383
4.3021 | -81.1378
4.3029 | -81.1349
4.3037 | -81.1329
4.3044 | -81.1331
4.3052 | -81.1323
4.306 | -81.1306
4.3068 | -81.1286
4.3075 | -81.1273
4.3083 | -81.126
4.3091 | -81.1229
4.3099 | -81.1212
4.3106 | -81.1243
4.3114 | -81.1251
4.3122 | -81.1242
4.313 | -81.1254
4.3137 | -81.125
4.3145 | -81.1276
4.3153 | -81.1255
4.3161 | -81.1229
4.3168 | -81.1216
4.3176 | -81.1201
4.3184 | -81.1209
4.3192 | -81.1223
4.3199 | -81.1201
4.3207 | -81.1168
4.3215 | -81.115
4.3223 | -81.1135
4.323 | -81.1123
4.3238 | -81.1119
4.3246 | -81.1105
4.3254 | -81.1093
4.3261 | -81.1104
4.3269 | -81.1093
4.3277 | -81.1105
4.3285 | -81.1159
4.3292 | -81.1166
4.33 | -81.1155
4.3308 | -81.1148
4.3316 | -81.1152
4.3323 | -81.1118
4.3331 | -81.1107
4.3339 | -81.1102
4.3347 | -81.109
4.3354 | -81.1085
4.3362 | -81.109
4.337 | -81.1076
4.3378 | -81.1047
4.3385 | -81.1009
4.3393 | -81.1001
4.3401 | -81.1013
4.3409 | -81.1004
4.3416 | -81.099
4.3424 | -81.0967
4.3432 | -81.0972
4.344 | -81.0982
4.3447 | -81.098
4.3455 | -81.0979
4.3463 | -81.0978
4.3471 | -81.0972
4.3478 | -81.0951
4.3486 | -81.093
4.3494 | -81.0936
4.3502 | -81.095
4.3509 | -81.0951
4.3517 | -81.0957
4.3525 | -81.0956
4.3532 | -81.0974
4.354 | -81.1012
4.3548 | -81.1063
4.3556 | -81.108
4.3563 | -81.1072
4.3571 | -81.1075
4.3579 | -81.1079
4.3587 | -81.1091
4.3594 | -81.1095
4.3602 | -81.111
4.361 | -81.1117
4.3618 | -81.1126
4.3625 | -81.1137
4.3633 | -81.1138
4.3641 | -81.1137
4.3649 | -81.1142
4.3656 | -81.1166
4.3664 | -81.1195
4.3672 | -81.1212
4.368 | -81.1204
4.3687 | -81.119
4.3695 | -81.1186
4.3703 | -81.1199
4.3711 | -81.122
4.3718 | -81.1235
4.3726 | -81.1245
4.3734 | -81.1241
4.3742 | -81.1242
4.3749 | -81.1257
4.3757 | -81.1286
4.3765 | -81.1315
4.3773 | -81.1331
4.378 | -81.135
4.3788 | -81.1366
4.3796 | -81.1394
4.3804 | -81.1386
4.3811 | -81.1446
4.3819 | -81.1464
4.3827 | -81.1502
4.3835 | -81.1554
4.3842 | -81.1567
4.385 | -81.1587
4.3858 | -81.1617
4.3866 | -81.1646
4.3873 | -81.1691
4.3881 | -81.1711
4.3889 | -81.1726
4.3897 | -81.1762
4.3904 | -81.1782
4.3912 | -81.1804
4.392 | -81.183
4.3928 | -81.1853
4.3935 | -81.1843
4.3943 | -81.1842
4.3951 | -81.1869
4.3959 | -81.1904
4.3966 | -81.1942
4.3974 | -81.1968
4.3982 | -81.1997
4.399 | -81.2044
4.3997 | -81.2115
4.4005 | -81.2145
4.4013 | -81.2174
4.4021 | -81.2201
4.4028 | -81.2228
4.4036 | -81.2249
4.4044 | -81.2269
4.4052 | -81.2315
4.4059 | -81.2381
4.4067 | -81.2425
4.4075 | -81.2462
4.4083 | -81.2536
4.409 | -81.2599
4.4098 | -81.2633
4.4106 | -81.2684
4.4114 | -81.2725
4.4121 | -81.2771
4.4129 | -81.2826
4.4137 | -81.2887
4.4145 | -81.2907
4.4152 | -81.2924
4.416 | -81.2964
4.4168 | -81.3004
4.4176 | -81.3037
4.4183 | -81.3061
4.4191 | -81.3087
4.4199 | -81.3118
4.4207 | -81.3157
4.4214 | -81.3199
4.4222 | -81.3248
4.423 | -81.3292
4.4238 | -81.3334
4.4245 | -81.3369
4.4253 | -81.3378
4.4261 | -81.3384
4.4268 | -81.3436
4.4276 | -81.3497
4.4284 | -81.3499
4.4292 | -81.349
4.4299 | -81.3476
4.4307 | -81.3498
4.4315 | -81.3534
4.4323 | -81.3581
4.433 | -81.3622
4.4338 | -81.366
4.4346 | -81.3703
4.4354 | -81.3744
4.4361 | -81.3772
4.4369 | -81.3799
4.4377 | -81.3842
4.4385 | -81.3908
4.4392 | -81.3949
4.44 | -81.3995
4.4408 | -81.4049
4.4416 | -81.4064
4.4423 | -81.4098
4.4431 | -81.4131
4.4439 | -81.4164
4.4447 | -81.4193
4.4454 | -81.4251
4.4462 | -81.4293
4.447 | -81.4322
4.4478 | -81.4327
4.4485 | -81.4351
4.4493 | -81.4386
4.4501 | -81.4403
4.4509 | -81.4426
4.4516 | -81.4465
4.4524 | -81.4497
4.4532 | -81.4528
4.454 | -81.4571
4.4547 | -81.4618
4.4555 | -81.4662
4.4563 | -81.468
4.4571 | -81.4698
4.4578 | -81.4704
4.4586 | -81.4715
4.4594 | -81.4734
4.4602 | -81.4792
4.4609 | -81.4843
4.4617 | -81.487
4.4625 | -81.492
4.4633 | -81.497
4.464 | -81.4993
4.4648 | -81.502
4.4656 | -81.5037
4.4664 | -81.5052
4.4671 | -81.5075
4.4679 | -81.5085
4.4687 | -81.5103
4.4695 | -81.5136
4.4702 | -81.5153
4.471 | -81.5163
4.4718 | -81.5141
4.4726 | -81.5157
4.4733 | -81.5192
4.4741 | -81.5198
4.4749 | -81.5229
4.4757 | -81.526
4.4764 | -81.5275
4.4772 | -81.5284
4.478 | -81.5299
4.4788 | -81.529
4.4795 | -81.5299
4.4803 | -81.5351
4.4811 | -81.5391
4.4819 | -81.5427
4.4826 | -81.5439
4.4834 | -81.5457
4.4842 | -81.5488
4.485 | -81.5511
4.4857 | -81.5523
4.4865 | -81.5536
4.4873 | -81.5551
4.4881 | -81.557
4.4888 | -81.5577
4.4896 | -81.558
4.4904 | -81.5589
4.4912 | -81.5601
4.4919 | -81.5608
4.4927 | -81.5613
4.4935 | -81.5601
4.4943 | -81.5566
4.495 | -81.5553
4.4958 | -81.5523
4.4966 | -81.5539
4.4973 | -81.5559
4.4981 | -81.5568
4.4989 | -81.5586
4.4997 | -81.5596
4.5004 | -81.5594
4.5012 | -81.5602
4.502 | -81.5597
4.5028 | -81.5609
4.5035 | -81.5629
4.5043 | -81.5647
4.5051 | -81.5647
4.5059 | -81.5653
4.5066 | -81.5654
4.5074 | -81.5656
4.5082 | -81.5661
4.509 | -81.5682
4.5097 | -81.5741
4.5105 | -81.5768
4.5113 | -81.5789
4.5121 | -81.5831
4.5128 | -81.5838
4.5136 | -81.583
4.5144 | -81.5842
4.5152 | -81.5816
4.5159 | -81.5817
4.5167 | -81.5806
4.5175 | -81.5813
4.5183 | -81.5825
4.519 | -81.5819
4.5198 | -81.5807
4.5206 | -81.5801
4.5214 | -81.5773
4.5221 | -81.5768
4.5229 | -81.5787
4.5237 | -81.5784
4.5245 | -81.5762
4.5252 | -81.5756
4.526 | -81.5752
4.5268 | -81.5697
4.5276 | -81.5686
4.5283 | -81.5686
4.5291 | -81.5689
4.5299 | -81.5691
4.5307 | -81.5688
4.5314 | -81.5743
4.5322 | -81.5753
4.533 | -81.5744
4.5338 | -81.5764
4.5345 | -81.5773
4.5353 | -81.576
4.5361 | -81.5748
4.5369 | -81.5771
4.5376 | -81.5784
4.5384 | -81.5782
4.5392 | -81.5791
4.54 | -81.5793
4.5407 | -81.5792
4.5415 | -81.579
4.5423 | -81.5776
4.5431 | -81.577
4.5438 | -81.5759
4.5446 | -81.5739
4.5454 | -81.5744
4.5462 | -81.5731
4.5469 | -81.57
4.5477 | -81.5687
4.5485 | -81.5689
4.5493 | -81.5695
4.55 | -81.5685
4.5508 | -81.5672
4.5516 | -81.5616
4.5524 | -81.5535
4.5531 | -81.5521
4.5539 | -81.5523
4.5547 | -81.5518
4.5555 | -81.5535
4.5562 | -81.5537
4.557 | -81.553
4.5578 | -81.5532
4.5586 | -81.5541
4.5593 | -81.5569
4.5601 | -81.5564
4.5609 | -81.5545
4.5617 | -81.5546
4.5624 | -81.5604
4.5632 | -81.5629
4.564 | -81.5647
4.5648 | -81.5647
4.5655 | -81.5638
4.5663 | -81.5642
4.5671 | -81.5632
4.5679 | -81.5631
4.5686 | -81.5643
4.5694 | -81.5647
4.5702 | -81.5645
4.5709 | -81.5602
4.5717 | -81.5614
4.5725 | -81.5636
4.5733 | -81.5628
4.574 | -81.5626
4.5748 | -81.5629
4.5756 | -81.5638
4.5764 | -81.5635
4.5771 | -81.5631
4.5779 | -81.5661
4.5787 | -81.5667
4.5795 | -81.5688
4.5802 | -81.5703
4.581 | -81.57
4.5818 | -81.5714
4.5826 | -81.5717
4.5833 | -81.5723
4.5841 | -81.5731
4.5849 | -81.5761
4.5857 | -81.5765
4.5864 | -81.5772
4.5872 | -81.5798
4.588 | -81.5798
4.5888 | -81.5818
4.5895 | -81.5829
4.5903 | -81.5828
4.5911 | -81.5805
4.5919 | -81.5801
4.5926 | -81.5814
4.5934 | -81.581
4.5942 | -81.5762
4.595 | -81.5732
4.5957 | -81.5707
4.5965 | -81.5685
4.5973 | -81.569
4.5981 | -81.5705
4.5988 | -81.5723
4.5996 | -81.5729
4.6004 | -81.5719
4.6012 | -81.572
4.6019 | -81.574
4.6027 | -81.5786
4.6035 | -81.5812
4.6043 | -81.5832
4.605 | -81.5846
4.6058 | -81.5871
4.6066 | -81.5894
4.6074 | -81.5926
4.6081 | -81.5952
4.6089 | -81.5974
4.6097 | -81.5973
4.6105 | -81.5945
4.6112 | -81.5948
4.612 | -81.5982
4.6128 | -81.5976
4.6136 | -81.5979
4.6143 | -81.5993
4.6151 | -81.5994
4.6159 | -81.5986
4.6167 | -81.6004
4.6174 | -81.6022
4.6182 | -81.6025
4.619 | -81.6035
4.6198 | -81.6059
4.6205 | -81.6067
4.6213 | -81.6068
4.6221 | -81.6081
4.6229 | -81.6082
4.6236 | -81.6084
4.6244 | -81.6067
4.6252 | -81.6051
4.626 | -81.6055
4.6267 | -81.6052
4.6275 | -81.6065
4.6283 | -81.6085
4.6291 | -81.6077
4.6298 | -81.6076
4.6306 | -81.6077
4.6314 | -81.6069
4.6322 | -81.6051
4.6329 | -81.6057
4.6337 | -81.6067
4.6345 | -81.6076
4.6353 | -81.6073
4.636 | -81.6018
4.6368 | -81.5952
4.6376 | -81.5931
4.6384 | -81.5929
4.6391 | -81.5919
4.6399 | -81.5902
4.6407 | -81.5888
4.6415 | -81.5882
4.6422 | -81.5862
4.643 | -81.5843
4.6438 | -81.5821
4.6445 | -81.5841
4.6453 | -81.5844
4.6461 | -81.5833
4.6469 | -81.5828
4.6476 | -81.5823
4.6484 | -81.5824
4.6492 | -81.5825
4.65 | -81.5823
4.6507 | -81.5821
4.6515 | -81.5844
4.6523 | -81.5875
4.6531 | -81.5865
4.6538 | -81.5852
4.6546 | -81.5837
4.6554 | -81.5836
4.6562 | -81.5845
4.6569 | -81.5842
4.6577 | -81.5814
4.6585 | -81.5788
4.6593 | -81.5776
4.66 | -81.5765
4.6608 | -81.5759
4.6616 | -81.5748
4.6624 | -81.5753
4.6631 | -81.5765
4.6639 | -81.577
4.6647 | -81.5757
4.6655 | -81.573
4.6662 | -81.5664
4.667 | -81.5617
4.6678 | -81.5599
4.6686 | -81.5582
4.6693 | -81.5557
4.6701 | -81.5528
4.6709 | -81.5502
4.6717 | -81.5493
4.6724 | -81.5478
4.6732 | -81.5481
4.674 | -81.5466
4.6748 | -81.546
4.6755 | -81.546
4.6763 | -81.5408
4.6771 | -81.5365
4.6779 | -81.5318
4.6786 | -81.529
4.6794 | -81.5275
4.6802 | -81.5276
4.681 | -81.524
4.6817 | -81.5198
4.6825 | -81.5174
4.6833 | -81.5165
4.6841 | -81.5169
4.6848 | -81.5169
4.6856 | -81.5134
4.6864 | -81.5074
4.6872 | -81.5029
4.6879 | -81.5009
4.6887 | -81.501
4.6895 | -81.5001
4.6903 | -81.4977
4.691 | -81.4942
4.6918 | -81.4896
4.6926 | -81.4846
4.6934 | -81.4822
4.6941 | -81.4796
4.6949 | -81.477
4.6957 | -81.4737
4.6965 | -81.4739
4.6972 | -81.471
4.698 | -81.4668
4.6988 | -81.4611
4.6996 | -81.4571
4.7003 | -81.4541
4.7011 | -81.4513
4.7019 | -81.4494
4.7027 | -81.4465
4.7034 | -81.443
4.7042 | -81.4402
4.705 | -81.4388
4.7058 | -81.4379
4.7065 | -81.4356
4.7073 | -81.4306
4.7081 | -81.4257
4.7089 | -81.4219
4.7096 | -81.4187
4.7104 | -81.419
4.7112 | -81.4178
4.712 | -81.4141
4.7127 | -81.4109
4.7135 | -81.4049
4.7143 | -81.4012
4.7151 | -81.3992
4.7158 | -81.3945
4.7166 | -81.3922
4.7174 | -81.3907
4.7182 | -81.3902
4.7189 | -81.3878
4.7197 | -81.3848
4.7205 | -81.3833
4.7212 | -81.3819
4.722 | -81.38
4.7228 | -81.3789
4.7236 | -81.3751
4.7243 | -81.3687
4.7251 | -81.3651
4.7259 | -81.3609
4.7267 | -81.3551
4.7274 | -81.3511
4.7282 | -81.3488
4.729 | -81.3449
4.7298 | -81.3426
4.7305 | -81.3398
4.7313 | -81.3369
4.7321 | -81.3343
4.7329 | -81.3321
4.7336 | -81.3295
4.7344 | -81.3268
4.7352 | -81.323
4.736 | -81.3199
4.7367 | -81.3179
4.7375 | -81.314
4.7383 | -81.3117
4.7391 | -81.3068
4.7398 | -81.2974
4.7406 | -81.2868
4.7414 | -81.2823
4.7422 | -81.2778
4.7429 | -81.2727
4.7437 | -81.2697
4.7445 | -81.2678
4.7453 | -81.2678
4.746 | -81.2659
4.7468 | -81.2628
4.7476 | -81.261
4.7484 | -81.2616
4.7491 | -81.2621
4.7499 | -81.2628
4.7507 | -81.2588
4.7515 | -81.2551
4.7522 | -81.2516
4.753 | -81.248
4.7538 | -81.2448
4.7546 | -81.241
4.7553 | -81.2378
4.7561 | -81.2361
4.7569 | -81.235
4.7577 | -81.2322
4.7584 | -81.2285
4.7592 | -81.2255
4.76 | -81.2221
4.7608 | -81.2191
4.7615 | -81.2154
4.7623 | -81.2137
4.7631 | -81.2094
4.7639 | -81.2076
4.7646 | -81.2065
4.7654 | -81.2035
4.7662 | -81.1999
4.767 | -81.1971
4.7677 | -81.194
4.7685 | -81.1924
4.7693 | -81.1913
4.7701 | -81.1893
4.7708 | -81.1858
4.7716 | -81.1816
4.7724 | -81.1779
4.7732 | -81.1748
4.7739 | -81.1719
4.7747 | -81.1688
4.7755 | -81.1657
4.7763 | -81.1635
4.777 | -81.1605
4.7778 | -81.157
4.7786 | -81.1543
4.7794 | -81.1535
4.7801 | -81.1527
4.7809 | -81.1513
4.7817 | -81.1488
4.7825 | -81.1474
4.7832 | -81.1477
4.784 | -81.1438
4.7848 | -81.1403
4.7856 | -81.1352
4.7863 | -81.1325
4.7871 | -81.1303
4.7879 | -81.1297
4.7887 | -81.1297
4.7894 | -81.1267
4.7902 | -81.1239
4.791 | -81.1179
4.7918 | -81.1117
4.7925 | -81.1097
4.7933 | -81.1076
4.7941 | -81.1049
4.7949 | -81.1059
4.7956 | -81.1046
4.7964 | -81.1003
4.7972 | -81.0961
4.7979 | -81.0927
4.7987 | -81.0891
4.7995 | -81.086
4.8003 | -81.0845
4.801 | -81.0825
4.8018 | -81.082
4.8026 | -81.0812
4.8034 | -81.0793
4.8041 | -81.0768
4.8049 | -81.0729
4.8057 | -81.07
4.8065 | -81.0685
4.8072 | -81.0658
4.808 | -81.0643
4.8088 | -81.062
4.8096 | -81.0591
4.8103 | -81.0563
4.8111 | -81.0538
4.8119 | -81.0515
4.8127 | -81.0458
4.8134 | -81.0434
4.8142 | -81.0412
4.815 | -81.0377
4.8158 | -81.0359
4.8165 | -81.0348
4.8173 | -81.0333
4.8181 | -81.0337
4.8189 | -81.0363
4.8196 | -81.0381
4.8204 | -81.0403
4.8212 | -81.0389
4.822 | -81.0366
4.8227 | -81.0324
4.8235 | -81.03
4.8243 | -81.0305
4.8251 | -81.0298
4.8258 | -81.0278
4.8266 | -81.0283
4.8274 | -81.0272
4.8282 | -81.0276
4.8289 | -81.026
4.8297 | -81.0229
4.8305 | -81.0224
4.8313 | -81.0226
4.832 | -81.0234
4.8328 | -81.0242
4.8336 | -81.024
4.8344 | -81.0233
4.8351 | -81.0229
4.8359 | -81.0227
4.8367 | -81.023
4.8375 | -81.0226
4.8382 | -81.0211
4.839 | -81.0166
4.8398 | -81.0118
4.8406 | -81.0121
4.8413 | -81.0121
4.8421 | -81.0111
4.8429 | -81.0117
4.8437 | -81.0112
4.8444 | -81.0098
4.8452 | -81.0095
4.846 | -81.0108
4.8468 | -81.0118
4.8475 | -81.0125
4.8483 | -81.011
4.8491 | -81.01
4.8499 | -81.0127
4.8506 | -81.0108
4.8514 | -81.0083
4.8522 | -81.0068
4.853 | -81.0061
4.8537 | -81.0102
4.8545 | -81.01
4.8553 | -81.0089
4.8561 | -81.0093
4.8568 | -81.0087
4.8576 | -81.0075
4.8584 | -81.0075
4.8592 | -81.0078
4.8599 | -81.0081
4.8607 | -81.0076
4.8615 | -81.008
4.8623 | -81.0075
4.863 | -81.0088
4.8638 | -81.0097
4.8646 | -81.0094
4.8654 | -81.0101
4.8661 | -81.0101
4.8669 | -81.0098
4.8677 | -81.0105
4.8685 | -81.0111
4.8692 | -81.0117
4.87 | -81.0141
4.8708 | -81.0132
4.8715 | -81.0122
4.8723 | -81.0132
4.8731 | -81.0138
4.8739 | -81.0147
4.8746 | -81.0174
4.8754 | -81.0193
4.8762 | -81.0207
4.877 | -81.0219
4.8777 | -81.0228
4.8785 | -81.0245
4.8793 | -81.0247
4.8801 | -81.0273
4.8808 | -81.0328
4.8816 | -81.0342
4.8824 | -81.0345
4.8832 | -81.0328
4.8839 | -81.0333
4.8847 | -81.035
4.8855 | -81.0354
4.8863 | -81.0375
4.887 | -81.0394
4.8878 | -81.0407
4.8886 | -81.0417
4.8894 | -81.0433
4.8901 | -81.044
4.8909 | -81.0451
4.8917 | -81.0463
4.8925 | -81.0485
4.8932 | -81.0492
4.894 | -81.0494
4.8948 | -81.052
4.8956 | -81.0516
4.8963 | -81.0538
4.8971 | -81.0571
4.8979 | -81.0602
4.8987 | -81.0632
4.8994 | -81.0626
4.9002 | -81.0639
4.901 | -81.0693
4.9018 | -81.0728
4.9025 | -81.0737
4.9033 | -81.0746
4.9041 | -81.0758
4.9049 | -81.0775
4.9056 | -81.0784
4.9064 | -81.0788
4.9072 | -81.0801
4.908 | -81.0791
4.9087 | -81.0765
4.9095 | -81.076
4.9103 | -81.0773
4.9111 | -81.079
4.9118 | -81.08
4.9126 | -81.0806
4.9134 | -81.0812
4.9142 | -81.0785
4.9149 | -81.0789
4.9157 | -81.0797
4.9165 | -81.0787
4.9173 | -81.0802
4.918 | -81.0802
4.9188 | -81.0787
4.9196 | -81.0809
4.9204 | -81.0822
4.9211 | -81.0834
4.9219 | -81.0858
4.9227 | -81.0896
4.9235 | -81.0902
4.9242 | -81.0899
4.925 | -81.0937
4.9258 | -81.1
4.9266 | -81.1048
4.9273 | -81.1083
4.9281 | -81.1117
4.9289 | -81.1113
4.9297 | -81.1108
4.9304 | -81.1106
4.9312 | -81.1107
4.932 | -81.1123
4.9328 | -81.1138
4.9335 | -81.1147
4.9343 | -81.1155
4.9351 | -81.1161
4.9359 | -81.1157
4.9366 | -81.1164
4.9374 | -81.1185
4.9382 | -81.1199
4.939 | -81.122
4.9397 | -81.1234
4.9405 | -81.1245
4.9413 | -81.1247
4.9421 | -81.1263
4.9428 | -81.13
4.9436 | -81.1345
4.9444 | -81.136
4.9451 | -81.136
4.9459 | -81.1379
4.9467 | -81.1411
4.9475 | -81.1428
4.9482 | -81.144
4.949 | -81.1453
4.9498 | -81.1458
4.9506 | -81.1477
4.9513 | -81.1494
4.9521 | -81.1496
4.9529 | -81.1497
4.9537 | -81.1517
4.9544 | -81.1529
4.9552 | -81.1531
4.956 | -81.1535
4.9568 | -81.1548
4.9575 | -81.1578
4.9583 | -81.1614
4.9591 | -81.1629
4.9599 | -81.1655
4.9606 | -81.1653
4.9614 | -81.1662
4.9622 | -81.1637
4.963 | -81.165
4.9637 | -81.1674
4.9645 | -81.1691
4.9653 | -81.169
4.9661 | -81.1693
4.9668 | -81.171
4.9676 | -81.1733
4.9684 | -81.1751
4.9692 | -81.1763
4.9699 | -81.178
4.9707 | -81.1836
4.9715 | -81.1903
4.9723 | -81.1924
4.973 | -81.1914
4.9738 | -81.191
4.9746 | -81.192
4.9754 | -81.1903
4.9761 | -81.1901
4.9769 | -81.1919
4.9777 | -81.1942
4.9785 | -81.1964
4.9792 | -81.1955
4.98 | -81.1942
4.9808 | -81.1955
4.9816 | -81.1959
4.9823 | -81.1967
4.9831 | -81.1969
4.9839 | -81.1977
4.9847 | -81.1987
4.9854 | -81.1978
4.9862 | -81.196
4.987 | -81.195
4.9878 | -81.196
4.9885 | -81.1967
4.9893 | -81.1975
4.9901 | -81.1981
4.9909 | -81.1999
4.9916 | -81.201
4.9924 | -81.202
4.9932 | -81.2029
4.994 | -81.2056
4.9947 | -81.207
4.9955 | -81.2079
4.9963 | -81.2093
4.9971 | -81.21
4.9978 | -81.2123
4.9986 | -81.2135
4.9994 | -81.2141
5.0002 | -81.2139
5.0009 | -81.2138
5.0017 | -81.2157
5.0025 | -81.2215
5.0033 | -81.2254
5.004 | -81.2267
5.0048 | -81.2273
5.0056 | -81.229
5.0064 | -81.2322
5.0071 | -81.235
5.0079 | -81.2372
5.0087 | -81.2387
5.0095 | -81.2393
5.0102 | -81.2405
5.011 | -81.2415
5.0118 | -81.2432
5.0126 | -81.2445
5.0133 | -81.2459
5.0141 | -81.2507
5.0149 | -81.2546
5.0156 | -81.2582
5.0164 | -81.2612
5.0172 | -81.2622
5.018 | -81.2629
5.0187 | -81.2659
5.0195 | -81.2689
5.0203 | -81.2711
5.0211 | -81.2739
5.0218 | -81.2746
5.0226 | -81.2751
5.0234 | -81.2789
5.0242 | -81.2833
5.0249 | -81.2854
5.0257 | -81.2866
5.0265 | -81.2886
5.0273 | -81.2907
5.028 | -81.2925
5.0288 | -81.296
5.0296 | -81.298
5.0304 | -81.3015
5.0311 | -81.3059
5.0319 | -81.3063
5.0327 | -81.3078
5.0335 | -81.3103
5.0342 | -81.3126
5.035 | -81.3161
5.0358 | -81.3184
5.0366 | -81.3208
5.0373 | -81.3233
5.0381 | -81.3279
5.0389 | -81.3311
5.0397 | -81.3337
5.0404 | -81.338
5.0412 | -81.3411
5.042 | -81.3439
5.0428 | -81.3468
5.0435 | -81.3493
5.0443 | -81.3518
5.0451 | -81.3536
5.0459 | -81.3567
5.0466 | -81.3624
5.0474 | -81.3669
5.0482 | -81.3691
5.049 | -81.3731
5.0497 | -81.3765
5.0505 | -81.3795
5.0513 | -81.3838
5.0521 | -81.3866
5.0528 | -81.3881
5.0536 | -81.394
5.0544 | -81.3972
5.0552 | -81.3997
5.0559 | -81.4032
5.0567 | -81.4081
5.0575 | -81.411
5.0583 | -81.4135
5.059 | -81.4163
5.0598 | -81.4185
5.0606 | -81.4218
5.0614 | -81.4252
5.0621 | -81.4289
5.0629 | -81.4318
5.0637 | -81.4342
5.0645 | -81.4386
5.0652 | -81.4447
5.066 | -81.4493
5.0668 | -81.4566
5.0676 | -81.4643
5.0683 | -81.469
5.0691 | -81.472
5.0699 | -81.476
5.0707 | -81.48
5.0714 | -81.4836
5.0722 | -81.4876
5.073 | -81.4898
5.0738 | -81.4905
5.0745 | -81.4951
5.0753 | -81.5
5.0761 | -81.5036
5.0769 | -81.5078
5.0776 | -81.5129
5.0784 | -81.5177
5.0792 | -81.5239
5.08 | -81.5295
5.0807 | -81.5338
5.0815 | -81.5413
5.0823 | -81.552
5.0831 | -81.5639
5.0838 | -81.5693
5.0846 | -81.5749
5.0854 | -81.5817
5.0862 | -81.5862
5.0869 | -81.5908
5.0877 | -81.5943
5.0885 | -81.5988
5.0892 | -81.604
5.09 | -81.609
5.0908 | -81.6152
5.0916 | -81.6216
5.0923 | -81.6277
5.0931 | -81.6328
5.0939 | -81.6368
5.0947 | -81.6418
5.0954 | -81.647
5.0962 | -81.6547
5.097 | -81.6581
5.0978 | -81.6618
5.0985 | -81.6676
5.0993 | -81.6729
5.1001 | -81.6771
5.1009 | -81.6828
5.1016 | -81.689
5.1024 | -81.6937
5.1032 | -81.6996
5.104 | -81.7063
5.1047 | -81.7137
5.1055 | -81.7208
5.1063 | -81.7256
5.1071 | -81.733
5.1078 | -81.7402
5.1086 | -81.7454
5.1094 | -81.7519
5.1102 | -81.7584
5.1109 | -81.7646
5.1117 | -81.7718
5.1125 | -81.7788
5.1133 | -81.7853
5.114 | -81.7923
5.1148 | -81.7999
5.1156 | -81.8096
5.1164 | -81.8188
5.1171 | -81.8258
5.1179 | -81.8328
5.1187 | -81.8427
5.1195 | -81.85
5.1202 | -81.8566
5.121 | -81.8618
5.1218 | -81.8687
5.1226 | -81.8778
5.1233 | -81.8861
5.1241 | -81.8935
5.1249 | -81.901
5.1257 | -81.9082
5.1264 | -81.916
5.1272 | -81.9235
5.128 | -81.931
5.1288 | -81.9378
5.1295 | -81.9462
5.1303 | -81.956
5.1311 | -81.9664
5.1319 | -81.972
5.1326 | -81.9794
5.1334 | -81.9878
5.1342 | -81.996
5.135 | -82.0054
5.1357 | -82.016
5.1365 | -82.0289
5.1373 | -82.0404
5.1381 | -82.0456
5.1388 | -82.0521
5.1396 | -82.0616
5.1404 | -82.07
5.1412 | -82.075
5.1419 | -82.0819
5.1427 | -82.0893
5.1435 | -82.097
5.1443 | -82.1047
5.145 | -82.1128
5.1458 | -82.121
5.1466 | -82.1251
5.1474 | -82.1305
5.1481 | -82.1377
5.1489 | -82.1455
5.1497 | -82.1521
5.1505 | -82.1594
5.1512 | -82.1699
5.152 | -82.1789
5.1528 | -82.1881
5.1536 | -82.1957
5.1543 | -82.2043
5.1551 | -82.2112
5.1559 | -82.2178
5.1567 | -82.2231
5.1574 | -82.2293
5.1582 | -82.2357
5.159 | -82.2428
5.1598 | -82.2554
5.1605 | -82.264
5.1613 | -82.271
5.1621 | -82.2801
5.1628 | -82.2876
5.1636 | -82.2951
5.1644 | -82.3014
5.1652 | -82.3081
5.1659 | -82.3146
5.1667 | -82.3211
5.1675 | -82.3273
5.1683 | -82.3339
5.169 | -82.3411
5.1698 | -82.3484
5.1706 | -82.3574
5.1714 | -82.3637
5.1721 | -82.369
5.1729 | -82.3754
5.1737 | -82.3831
5.1745 | -82.3895
5.1752 | -82.3965
5.176 | -82.3991
5.1768 | -82.3983
5.1776 | -82.4041
5.1783 | -82.4131
5.1791 | -82.4203
5.1799 | -82.4263
5.1807 | -82.4321
5.1814 | -82.4373
5.1822 | -82.4435
5.183 | -82.4502
5.1838 | -82.4559
5.1845 | -82.462
5.1853 | -82.4668
5.1861 | -82.4724
5.1869 | -82.4782
5.1876 | -82.4857
5.1884 | -82.4961
5.1892 | -82.505
5.19 | -82.5078
5.1907 | -82.5139
5.1915 | -82.5186
5.1923 | -82.5228
5.1931 | -82.527
5.1938 | -82.5324
5.1946 | -82.5368
5.1954 | -82.5415
5.1962 | -82.5463
5.1969 | -82.5519
5.1977 | -82.5582
5.1985 | -82.5646
5.1993 | -82.5694
5.2 | -82.5737
5.2008 | -82.5792
5.2016 | -82.5872
5.2024 | -82.5937
5.2031 | -82.5993
5.2039 | -82.6039
5.2047 | -82.6099
5.2055 | -82.6138
5.2062 | -82.6162
5.207 | -82.6201
5.2078 | -82.6214
5.2086 | -82.6268
5.2093 | -82.6301
5.2101 | -82.6331
5.2109 | -82.6354
5.2117 | -82.638
5.2124 | -82.6413
5.2132 | -82.6468
5.214 | -82.6509
5.2148 | -82.6528
5.2155 | -82.6559
5.2163 | -82.6597
5.2171 | -82.6644
5.2179 | -82.6688
5.2186 | -82.6708
5.2194 | -82.6734
5.2202 | -82.6775
5.221 | -82.6817
5.2217 | -82.688
5.2225 | -82.6931
5.2233 | -82.6953
5.2241 | -82.6985
5.2248 | -82.7013
5.2256 | -82.7032
5.2264 | -82.7061
5.2272 | -82.7098
5.2279 | -82.712
5.2287 | -82.7135
5.2295 | -82.7135
5.2303 | -82.7111
5.231 | -82.7126
5.2318 | -82.7157
5.2326 | -82.7209
5.2334 | -82.7242
5.2341 | -82.7259
5.2349 | -82.7254
5.2357 | -82.725
5.2364 | -82.7246
5.2372 | -82.7267
5.238 | -82.7266
5.2388 | -82.726
5.2395 | -82.7266
5.2403 | -82.7253
5.2411 | -82.724
5.2419 | -82.7245
5.2426 | -82.7274
5.2434 | -82.7268
5.2442 | -82.7247
5.245 | -82.7241
5.2457 | -82.7258
5.2465 | -82.728
5.2473 | -82.7275
5.2481 | -82.7295
5.2488 | -82.7321
5.2496 | -82.7319
5.2504 | -82.7322
5.2512 | -82.7329
5.2519 | -82.733
5.2527 | -82.7321
5.2535 | -82.7347
5.2543 | -82.7343
5.255 | -82.7325
5.2558 | -82.7317
5.2566 | -82.7321
5.2574 | -82.7327
5.2581 | -82.7334
5.2589 | -82.7322
5.2597 | -82.7307
5.2605 | -82.7322
5.2612 | -82.7327
5.262 | -82.7349
5.2628 | -82.7386
5.2636 | -82.7393
5.2643 | -82.7396
5.2651 | -82.7398
5.2659 | -82.7377
5.2667 | -82.7372
5.2674 | -82.7372
5.2682 | -82.7406
5.269 | -82.7398
5.2698 | -82.739
5.2705 | -82.7408
5.2713 | -82.7408
5.2721 | -82.7398
5.2729 | -82.741
5.2736 | -82.7439
5.2744 | -82.7414
5.2752 | -82.74
5.276 | -82.7381
5.2767 | -82.7362
5.2775 | -82.7327
5.2783 | -82.7299
5.2791 | -82.7299
5.2798 | -82.7316
5.2806 | -82.7301
5.2814 | -82.7281
5.2822 | -82.7259
5.2829 | -82.7248
5.2837 | -82.725
5.2845 | -82.7254
5.2853 | -82.7251
5.286 | -82.7252
5.2868 | -82.7252
5.2876 | -82.725
5.2884 | -82.7241
5.2891 | -82.7217
5.2899 | -82.7191
5.2907 | -82.715
5.2915 | -82.7141
5.2922 | -82.7143
5.293 | -82.7127
5.2938 | -82.7116
5.2946 | -82.7106
5.2953 | -82.7102
5.2961 | -82.7103
5.2969 | -82.7095
5.2977 | -82.7098
5.2984 | -82.7083
5.2992 | -82.7066
5.3 | -82.7057
5.3008 | -82.7045
5.3015 | -82.7048
5.3023 | -82.7057
5.3031 | -82.7047
5.3039 | -82.7026
5.3046 | -82.7018
5.3054 | -82.7003
5.3062 | -82.6989
5.307 | -82.6946
5.3077 | -82.691
5.3085 | -82.6903
5.3093 | -82.6912
5.3101 | -82.6918
5.3108 | -82.6913
5.3116 | -82.6914
5.3124 | -82.6921
5.3131 | -82.6932
5.3139 | -82.6932
5.3147 | -82.6937
5.3155 | -82.693
5.3162 | -82.693
5.317 | -82.6918
5.3178 | -82.6933
5.3186 | -82.6939
5.3193 | -82.6969
5.3201 | -82.7
5.3209 | -82.704
5.3217 | -82.7044
5.3224 | -82.7044
5.3232 | -82.7044
5.324 | -82.7068
5.3248 | -82.7104
5.3255 | -82.7129
5.3263 | -82.7141
5.3271 | -82.7145
5.3279 | -82.7142
5.3286 | -82.715
5.3294 | -82.7167
5.3302 | -82.718
5.331 | -82.7194
5.3317 | -82.7227
5.3325 | -82.7251
5.3333 | -82.7263
5.3341 | -82.7318
5.3348 | -82.7348
5.3356 | -82.735
5.3364 | -82.7351
5.3372 | -82.7347
5.3379 | -82.7382
5.3387 | -82.7422
5.3395 | -82.7433
5.3403 | -82.745
5.341 | -82.7465
5.3418 | -82.7496
5.3426 | -82.7512
5.3434 | -82.7524
5.3441 | -82.7543
5.3449 | -82.7563
5.3457 | -82.7586
5.3465 | -82.7644
5.3472 | -82.769
5.348 | -82.7708
5.3488 | -82.7726
5.3496 | -82.7754
5.3503 | -82.7784
5.3511 | -82.7804
5.3519 | -82.7819
5.3527 | -82.7844
5.3534 | -82.7875
5.3542 | -82.7888
5.355 | -82.7913
5.3558 | -82.794
5.3565 | -82.7957
5.3573 | -82.7985
5.3581 | -82.8004
5.3589 | -82.803
5.3596 | -82.8066
5.3604 | -82.8107
5.3612 | -82.8147
5.362 | -82.8169
5.3627 | -82.8199
5.3635 | -82.8222
5.3643 | -82.8283
5.3651 | -82.8305
5.3658 | -82.8347
5.3666 | -82.8376
5.3674 | -82.8394
5.3682 | -82.8422
5.3689 | -82.8451
5.3697 | -82.8483
5.3705 | -82.8513
5.3713 | -82.8556
5.372 | -82.8587
5.3728 | -82.8635
5.3736 | -82.8677
5.3744 | -82.8714
5.3751 | -82.8758
5.3759 | -82.8804
5.3767 | -82.8847
5.3775 | -82.8896
5.3782 | -82.8938
5.379 | -82.8998
5.3798 | -82.9043
5.3806 | -82.9084
5.3813 | -82.9141
5.3821 | -82.9192
5.3829 | -82.9229
5.3837 | -82.9268
5.3844 | -82.9326
5.3852 | -82.938
5.386 | -82.9438
5.3867 | -82.9518
5.3875 | -82.9583
5.3883 | -82.962
5.3891 | -82.9659
5.3898 | -82.9709
5.3906 | -82.9752
5.3914 | -82.9779
5.3922 | -82.9811
5.3929 | -82.9864
5.3937 | -82.9926
5.3945 | -82.9989
5.3953 | -83.0062
5.396 | -83.0106
5.3968 | -83.0138
5.3976 | -83.0188
5.3984 | -83.0234
5.3991 | -83.0282
5.3999 | -83.034
5.4007 | -83.0407
5.4015 | -83.0468
5.4022 | -83.0517
5.403 | -83.056
5.4038 | -83.0601
5.4046 | -83.0648
5.4053 | -83.0696
5.4061 | -83.0751
5.4069 | -83.0808
5.4077 | -83.0868
5.4084 | -83.0925
5.4092 | -83.0984
5.41 | -83.1031
5.4108 | -83.1069
5.4115 | -83.1138
5.4123 | -83.1221
5.4131 | -83.128
5.4139 | -83.1342
5.4146 | -83.1399
5.4154 | -83.1462
5.4162 | -83.1523
5.417 | -83.1577
5.4177 | -83.1643
5.4185 | -83.1697
5.4193 | -83.1778
5.4201 | -83.1842
5.4208 | -83.1906
5.4216 | -83.1966
5.4224 | -83.2032
5.4232 | -83.21
5.4239 | -83.2171
5.4247 | -83.2248
5.4255 | -83.2305
5.4263 | -83.2357
5.427 | -83.2406
5.4278 | -83.2477
5.4286 | -83.2553
5.4294 | -83.262
5.4301 | -83.2686
5.4309 | -83.2726
5.4317 | -83.2793
5.4325 | -83.2854
5.4332 | -83.2916
5.434 | -83.2978
5.4348 | -83.3046
5.4356 | -83.3108
5.4363 | -83.3162
5.4371 | -83.3234
5.4379 | -83.33
5.4387 | -83.3391
5.4394 | -83.3509
5.4402 | -83.3593
5.441 | -83.3661
5.4418 | -83.3728
5.4425 | -83.3806
5.4433 | -83.387
5.4441 | -83.3916
5.4449 | -83.3927
5.4456 | -83.3967
5.4464 | -83.4016
5.4472 | -83.4075
5.448 | -83.4137
5.4487 | -83.4196
5.4495 | -83.4239
5.4503 | -83.4293
5.4511 | -83.4357
5.4518 | -83.4392
5.4526 | -83.4444
5.4534 | -83.4487
5.4542 | -83.4554
5.4549 | -83.464
5.4557 | -83.4729
5.4565 | -83.4783
5.4573 | -83.4817
5.458 | -83.4869
5.4588 | -83.4924
5.4596 | -83.4978
5.4603 | -83.5059
5.4611 | -83.5136
5.4619 | -83.516
5.4627 | -83.5203
5.4634 | -83.5278
5.4642 | -83.5306
5.465 | -83.5352
5.4658 | -83.542
5.4665 | -83.5471
5.4673 | -83.5531
5.4681 | -83.561
5.4689 | -83.5666
5.4696 | -83.5701
5.4704 | -83.5739
5.4712 | -83.5795
5.472 | -83.5842
5.4727 | -83.5879
5.4735 | -83.5925
5.4743 | -83.6004
5.4751 | -83.6057
5.4758 | -83.6093
5.4766 | -83.614
5.4774 | -83.62
5.4782 | -83.6238
5.4789 | -83.6279
5.4797 | -83.6328
5.4805 | -83.6373
5.4813 | -83.6418
5.482 | -83.6456
5.4828 | -83.6494
5.4836 | -83.6534
5.4844 | -83.6576
5.4851 | -83.6627
5.4859 | -83.6673
5.4867 | -83.6715
5.4875 | -83.6744
5.4882 | -83.6764
5.489 | -83.6808
5.4898 | -83.6864
5.4906 | -83.6932
5.4913 | -83.6988
5.4921 | -83.7033
5.4929 | -83.708
5.4937 | -83.7116
5.4944 | -83.7155
5.4952 | -83.7193
5.496 | -83.7228
5.4968 | -83.7257
5.4975 | -83.7299
5.4983 | -83.7301
5.4991 | -83.7347
5.4999 | -83.7393
5.5006 | -83.7427
5.5014 | -83.7441
5.5022 | -83.75
5.503 | -83.7535
5.5037 | -83.7556
5.5045 | -83.758
5.5053 | -83.7607
5.5061 | -83.7631
5.5068 | -83.7646
5.5076 | -83.7675
5.5084 | -83.774
5.5092 | -83.7788
5.5099 | -83.782
5.5107 | -83.7836
5.5115 | -83.786
5.5123 | -83.7924
5.513 | -83.7966
5.5138 | -83.8036
5.5146 | -83.8104
5.5154 | -83.8149
5.5161 | -83.8188
5.5169 | -83.822
5.5177 | -83.8283
5.5185 | -83.8323
5.5192 | -83.8353
5.52 | -83.8367
5.5208 | -83.8344
5.5216 | -83.8364
5.5223 | -83.8392
5.5231 | -83.8399
5.5239 | -83.8425
5.5247 | -83.8441
5.5254 | -83.8448
5.5262 | -83.845
5.527 | -83.8452
5.5278 | -83.848
5.5285 | -83.8542
5.5293 | -83.8584
5.5301 | -83.862
5.5309 | -83.8646
5.5316 | -83.865
5.5324 | -83.864
5.5332 | -83.8662
5.5339 | -83.8661
5.5347 | -83.8665
5.5355 | -83.8667
5.5363 | -83.8673
5.537 | -83.8667
5.5378 | -83.8681
5.5386 | -83.8688
5.5394 | -83.8695
5.5401 | -83.8706
5.5409 | -83.8712
5.5417 | -83.8713
5.5425 | -83.873
5.5432 | -83.8725
5.544 | -83.8746
5.5448 | -83.8773
5.5456 | -83.8786
5.5463 | -83.8793
5.5471 | -83.8797
5.5479 | -83.8788
5.5487 | -83.88
5.5494 | -83.8855
5.5502 | -83.8935
5.551 | -83.8955
5.5518 | -83.8955
5.5525 | -83.8952
5.5533 | -83.8948
5.5541 | -83.8947
5.5549 | -83.8963
5.5556 | -83.897
5.5564 | -83.8966
5.5572 | -83.8954
5.558 | -83.895
5.5587 | -83.8938
5.5595 | -83.8931
5.5603 | -83.8928
5.5611 | -83.8923
5.5618 | -83.8907
5.5626 | -83.8905
5.5634 | -83.8915
5.5642 | -83.8934
5.5649 | -83.8935
5.5657 | -83.8946
5.5665 | -83.8956
5.5673 | -83.8959
5.568 | -83.8965
5.5688 | -83.8974
5.5696 | -83.8975
5.5704 | -83.896
5.5711 | -83.8952
5.5719 | -83.8923
5.5727 | -83.8908
5.5735 | -83.891
5.5742 | -83.8899
5.575 | -83.8893
5.5758 | -83.8899
5.5766 | -83.8895
5.5773 | -83.8882
5.5781 | -83.887
5.5789 | -83.8849
5.5797 | -83.8824
5.5804 | -83.8825
5.5812 | -83.8832
5.582 | -83.8822
5.5828 | -83.8827
5.5835 | -83.8826
5.5843 | -83.8814
5.5851 | -83.8802
5.5859 | -83.8791
5.5866 | -83.8775
5.5874 | -83.8779
5.5882 | -83.8792
5.589 | -83.8788
5.5897 | -83.8782
5.5905 | -83.8776
5.5913 | -83.8771
5.5921 | -83.8764
5.5928 | -83.8763
5.5936 | -83.8755
5.5944 | -83.8736
5.5952 | -83.8719
5.5959 | -83.8704
5.5967 | -83.8696
5.5975 | -83.8693
5.5983 | -83.8686
5.599 | -83.8684
5.5998 | -83.869
5.6006 | -83.8708
5.6014 | -83.8713
5.6021 | -83.8712
5.6029 | -83.8712
5.6037 | -83.8714
5.6044 | -83.8711
5.6052 | -83.8706
5.606 | -83.8709
5.6068 | -83.8729
5.6075 | -83.8733
5.6083 | -83.8738
5.6091 | -83.8746
5.6099 | -83.8737
5.6106 | -83.8751
5.6114 | -83.8756
5.6122 | -83.8753
5.613 | -83.8783
5.6137 | -83.8847
5.6145 | -83.887
5.6153 | -83.8871
5.6161 | -83.89
5.6168 | -83.8919
5.6176 | -83.8922
5.6184 | -83.8936
5.6192 | -83.8959
5.6199 | -83.8968
5.6207 | -83.8971
5.6215 | -83.9006
5.6223 | -83.9034
5.623 | -83.9047
5.6238 | -83.9048
5.6246 | -83.9049
5.6254 | -83.9024
5.6261 | -83.9019
5.6269 | -83.9026
5.6277 | -83.9031
5.6285 | -83.9035
5.6292 | -83.9038
5.63 | -83.903
5.6308 | -83.9043
5.6316 | -83.9057
5.6323 | -83.9097
5.6331 | -83.9102
5.6339 | -83.9105
5.6347 | -83.9119
5.6354 | -83.9129
5.6362 | -83.9137
5.637 | -83.9156
5.6378 | -83.917
5.6385 | -83.9201
5.6393 | -83.923
5.6401 | -83.9235
5.6409 | -83.9254
5.6416 | -83.9274
5.6424 | -83.9293
5.6432 | -83.9325
5.644 | -83.9347
5.6447 | -83.9366
5.6455 | -83.94
5.6463 | -83.9414
5.6471 | -83.9427
5.6478 | -83.945
5.6486 | -83.9474
5.6494 | -83.9486
5.6502 | -83.9503
5.6509 | -83.9518
5.6517 | -83.9529
5.6525 | -83.954
5.6533 | -83.9551
5.654 | -83.9585
5.6548 | -83.9627
5.6556 | -83.9661
5.6564 | -83.9682
5.6571 | -83.9702
5.6579 | -83.9727
5.6587 | -83.975
5.6595 | -83.9794
5.6602 | -83.9836
5.661 | -83.9857
5.6618 | -83.9884
5.6626 | -83.9918
5.6633 | -83.9945
5.6641 | -83.9963
5.6649 | -83.9986
5.6657 | -84.002
5.6664 | -84.0058
5.6672 | -84.0083
5.668 | -84.0103
5.6688 | -84.0128
5.6695 | -84.0156
5.6703 | -84.0196
5.6711 | -84.0241
5.6719 | -84.0311
5.6726 | -84.0348
5.6734 | -84.036
5.6742 | -84.0395
5.675 | -84.0421
5.6757 | -84.0446
5.6765 | -84.0467
5.6773 | -84.0512
5.678 | -84.0543
5.6788 | -84.0573
5.6796 | -84.06
5.6804 | -84.065
5.6811 | -84.0705
5.6819 | -84.074
5.6827 | -84.0757
5.6835 | -84.0791
5.6842 | -84.082
5.685 | -84.0851
5.6858 | -84.0882
5.6866 | -84.0901
5.6873 | -84.0936
5.6881 | -84.0984
5.6889 | -84.103
5.6897 | -84.1116
5.6904 | -84.1179
5.6912 | -84.1197
5.692 | -84.1213
5.6928 | -84.1249
5.6935 | -84.1285
5.6943 | -84.1315
5.6951 | -84.1349
5.6959 | -84.1384
5.6966 | -84.1414
5.6974 | -84.1421
5.6982 | -84.1431
5.699 | -84.1443
5.6997 | -84.1488
5.7005 | -84.153
5.7013 | -84.1569
5.7021 | -84.16
5.7028 | -84.1638
5.7036 | -84.1665
5.7044 | -84.17
5.7052 | -84.1725
5.7059 | -84.1737
5.7067 | -84.1766
5.7075 | -84.1806
5.7083 | -84.1837
5.709 | -84.1873
5.7098 | -84.1905
5.7106 | -84.1931
5.7114 | -84.1964
5.7121 | -84.1994
5.7129 | -84.2033
5.7137 | -84.2058
5.7145 | -84.2063
5.7152 | -84.2056
5.716 | -84.2076
5.7168 | -84.213
5.7176 | -84.2172
5.7183 | -84.219
5.7191 | -84.2213
5.7199 | -84.2262
5.7207 | -84.2274
5.7214 | -84.2288
5.7222 | -84.2327
5.723 | -84.2362
5.7238 | -84.2381
5.7245 | -84.2413
5.7253 | -84.2446
5.7261 | -84.247
5.7269 | -84.2502
5.7276 | -84.2556
5.7284 | -84.2589
5.7292 | -84.2625
5.73 | -84.2653
5.7307 | -84.2691
5.7315 | -84.2741
5.7323 | -84.2759
5.7331 | -84.2786
5.7338 | -84.2816
5.7346 | -84.2846
5.7354 | -84.2861
5.7362 | -84.2861
5.7369 | -84.2885
5.7377 | -84.2908
5.7385 | -84.2941
5.7393 | -84.2973
5.74 | -84.3012
5.7408 | -84.3063
5.7416 | -84.309
5.7424 | -84.3094
5.7431 | -84.3106
5.7439 | -84.3155
5.7447 | -84.3178
5.7455 | -84.3192
5.7462 | -84.3213
5.747 | -84.3258
5.7478 | -84.3291
5.7486 | -84.3315
5.7493 | -84.3337
5.7501 | -84.336
5.7509 | -84.3376
5.7516 | -84.3403
5.7524 | -84.3426
5.7532 | -84.344
5.754 | -84.3469
5.7547 | -84.3515
5.7555 | -84.3537
5.7563 | -84.3565
5.7571 | -84.359
5.7578 | -84.3613
5.7586 | -84.3638
5.7594 | -84.3671
5.7602 | -84.3697
5.7609 | -84.3721
5.7617 | -84.3732
5.7625 | -84.3754
5.7633 | -84.3781
5.764 | -84.3803
5.7648 | -84.3809
5.7656 | -84.3815
5.7664 | -84.3795
5.7671 | -84.3812
5.7679 | -84.3845
5.7687 | -84.3855
5.7695 | -84.3855
5.7702 | -84.3872
5.771 | -84.3894
5.7718 | -84.3916
5.7726 | -84.3923
5.7733 | -84.3931
5.7741 | -84.3947
5.7749 | -84.3965
5.7757 | -84.3982
5.7764 | -84.4006
5.7772 | -84.4029
5.778 | -84.4038
5.7788 | -84.4062
5.7795 | -84.4078
5.7803 | -84.4126
5.7811 | -84.4172
5.7819 | -84.4196
5.7826 | -84.4202
5.7834 | -84.4214
5.7842 | -84.4232
5.785 | -84.4268
5.7857 | -84.4309
5.7865 | -84.4308
5.7873 | -84.429
5.7881 | -84.4304
5.7888 | -84.4311
5.7896 | -84.4317
5.7904 | -84.4329
5.7912 | -84.4339
5.7919 | -84.4352
5.7927 | -84.4364
5.7935 | -84.4385
5.7943 | -84.4408
5.795 | -84.4431
5.7958 | -84.4443
5.7966 | -84.4454
5.7974 | -84.446
5.7981 | -84.4475
5.7989 | -84.4506
5.7997 | -84.4511
5.8005 | -84.4531
5.8012 | -84.4564
5.802 | -84.4576
5.8028 | -84.4583
5.8036 | -84.4589
5.8043 | -84.4594
5.8051 | -84.4626
5.8059 | -84.4664
5.8067 | -84.4668
5.8074 | -84.4697
5.8082 | -84.471
5.809 | -84.4696
5.8098 | -84.4696
5.8105 | -84.4703
5.8113 | -84.4705
5.8121 | -84.4686
5.8129 | -84.4689
5.8136 | -84.4694
5.8144 | -84.4699
5.8152 | -84.4692
5.816 | -84.4694
5.8167 | -84.4705
5.8175 | -84.4725
5.8183 | -84.4743
5.8191 | -84.4741
5.8198 | -84.4739
5.8206 | -84.4724
5.8214 | -84.4731
5.8222 | -84.4753
5.8229 | -84.4734
5.8237 | -84.4728
5.8245 | -84.4726
5.8252 | -84.4725
5.826 | -84.4727
5.8268 | -84.4733
5.8276 | -84.475
5.8283 | -84.4734
5.8291 | -84.4702
5.8299 | -84.4701
5.8307 | -84.4698
5.8314 | -84.4713
5.8322 | -84.4721
5.833 | -84.4716
5.8338 | -84.4709
5.8345 | -84.4713
5.8353 | -84.4689
5.8361 | -84.4684
5.8369 | -84.4697
5.8376 | -84.4751
5.8384 | -84.4773
5.8392 | -84.4783
5.84 | -84.4787
5.8407 | -84.4767
5.8415 | -84.474
5.8423 | -84.473
5.8431 | -84.4733
5.8438 | -84.4727
5.8446 | -84.4746
5.8454 | -84.476
5.8462 | -84.4764
5.8469 | -84.4767
5.8477 | -84.4758
5.8485 | -84.4752
5.8493 | -84.4744
5.85 | -84.4724
5.8508 | -84.4737
5.8516 | -84.4736
5.8524 | -84.4729
5.8531 | -84.473
5.8539 | -84.472
5.8547 | -84.4719
5.8555 | -84.4756
5.8562 | -84.4758
5.857 | -84.4777
5.8578 | -84.477
5.8586 | -84.4794
5.8593 | -84.4831
5.8601 | -84.4843
5.8609 | -84.484
5.8617 | -84.4832
5.8624 | -84.4814
5.8632 | -84.4793
5.864 | -84.4804
5.8648 | -84.4828
5.8655 | -84.4808
5.8663 | -84.4775
5.8671 | -84.4756
5.8679 | -84.4749
5.8686 | -84.473
5.8694 | -84.471
5.8702 | -84.4693
5.871 | -84.468
5.8717 | -84.4653
5.8725 | -84.4627
5.8733 | -84.4633
5.8741 | -84.4636
5.8748 | -84.4623
5.8756 | -84.4613
5.8764 | -84.4597
5.8772 | -84.4598
5.8779 | -84.4593
5.8787 | -84.4577
5.8795 | -84.4578
5.8803 | -84.4569
5.881 | -84.4552
5.8818 | -84.4541
5.8826 | -84.4535
5.8834 | -84.4543
5.8841 | -84.4558
5.8849 | -84.4518
5.8857 | -84.4463
5.8865 | -84.4411
5.8872 | -84.439
5.888 | -84.4371
5.8888 | -84.4338
5.8896 | -84.4327
5.8903 | -84.432
5.8911 | -84.4319
5.8919 | -84.432
5.8927 | -84.4299
5.8934 | -84.4241
5.8942 | -84.422
5.895 | -84.4221
5.8958 | -84.4201
5.8965 | -84.4176
5.8973 | -84.4163
5.8981 | -84.4185
5.8989 | -84.4189
5.8996 | -84.418
5.9004 | -84.4174
5.9012 | -84.4175
5.9019 | -84.4167
5.9027 | -84.4158
5.9035 | -84.4139
5.9043 | -84.4144
5.905 | -84.4152
5.9058 | -84.4148
5.9066 | -84.4139
5.9074 | -84.4127
5.9081 | -84.4133
5.9089 | -84.4144
5.9097 | -84.4141
5.9105 | -84.4139
5.9112 | -84.4128
5.912 | -84.4121
5.9128 | -84.4103
5.9136 | -84.409
5.9143 | -84.4077
5.9151 | -84.4077
5.9159 | -84.4082
5.9167 | -84.4068
5.9174 | -84.4055
5.9182 | -84.406
5.919 | -84.4077
5.9198 | -84.4074
5.9205 | -84.4068
5.9213 | -84.4071
5.9221 | -84.4067
5.9229 | -84.406
5.9236 | -84.4048
5.9244 | -84.4033
5.9252 | -84.4016
5.926 | -84.4021
5.9267 | -84.4021
5.9275 | -84.4005
5.9283 | -84.4002
5.9291 | -84.3999
5.9298 | -84.3988
5.9306 | -84.3992
5.9314 | -84.4019
5.9322 | -84.4016
5.9329 | -84.4026
5.9337 | -84.4028
5.9345 | -84.4027
5.9353 | -84.403
5.936 | -84.4046
5.9368 | -84.403
5.9376 | -84.4031
5.9384 | -84.4024
5.9391 | -84.4018
5.9399 | -84.4008
5.9407 | -84.4005
5.9415 | -84.4003
5.9422 | -84.4008
5.943 | -84.4008
5.9438 | -84.3982
5.9446 | -84.3966
5.9453 | -84.3984
5.9461 | -84.3975
5.9469 | -84.3967
5.9477 | -84.396
5.9484 | -84.3969
5.9492 | -84.3967
5.95 | -84.3977
5.9508 | -84.3988
5.9515 | -84.398
5.9523 | -84.3975
5.9531 | -84.396
5.9539 | -84.3949
5.9546 | -84.3948
5.9554 | -84.3952
5.9562 | -84.3957
5.957 | -84.3956
5.9577 | -84.3967
5.9585 | -84.3963
5.9593 | -84.3948
5.9601 | -84.3949
5.9608 | -84.3948
5.9616 | -84.3946
5.9624 | -84.3957
5.9632 | -84.3966
5.9639 | -84.3967
5.9647 | -84.3959
5.9655 | -84.3962
5.9663 | -84.3964
5.967 | -84.3954
5.9678 | -84.3955
5.9686 | -84.3954
5.9694 | -84.3962
5.9701 | -84.397
5.9709 | -84.3969
5.9717 | -84.3978
5.9725 | -84.3993
5.9732 | -84.4007
5.974 | -84.4026
5.9748 | -84.4014
5.9755 | -84.3994
5.9763 | -84.3983
5.9771 | -84.3993
5.9779 | -84.4009
5.9786 | -84.4032
5.9794 | -84.4043
5.9802 | -84.4049
5.981 | -84.4062
5.9817 | -84.4085
5.9825 | -84.4088
5.9833 | -84.4074
5.9841 | -84.4091
5.9848 | -84.4097
5.9856 | -84.41
5.9864 | -84.4108
5.9872 | -84.4125
5.9879 | -84.4148
5.9887 | -84.4179
5.9895 | -84.4185
5.9903 | -84.4198
5.991 | -84.4201
5.9918 | -84.4203
5.9926 | -84.4217
5.9934 | -84.4207
5.9941 | -84.4201
5.9949 | -84.4204
5.9957 | -84.4221
5.9965 | -84.4223
5.9972 | -84.422
5.998 | -84.4218
5.9988 | -84.4233
5.9996 | -84.4242
6.0003 | -84.425
6.0011 | -84.4252
6.0019 | -84.4261
6.0027 | -84.4268
6.0034 | -84.4273
6.0042 | -84.4278
6.005 | -84.4274
6.0058 | -84.4266
6.0065 | -84.4266
6.0073 | -84.4268
6.0081 | -84.4274
6.0089 | -84.429
6.0096 | -84.4314
6.0104 | -84.4331
6.0112 | -84.4344
6.012 | -84.4353
6.0127 | -84.4375
6.0135 | -84.4379
6.0143 | -84.4408
6.0151 | -84.4426
6.0158 | -84.4429
6.0166 | -84.4437
6.0174 | -84.445
6.0182 | -84.4444
6.0189 | -84.4452
6.0197 | -84.4465
6.0205 | -84.4457
6.0213 | -84.4435
6.022 | -84.4444
6.0228 | -84.4444
6.0236 | -84.4461
6.0244 | -84.448
6.0251 | -84.449
6.0259 | -84.4506
6.0267 | -84.4513
6.0275 | -84.4514
6.0282 | -84.4506
6.029 | -84.4524
6.0298 | -84.4539
6.0306 | -84.4577
6.0313 | -84.4591
6.0321 | -84.4601
6.0329 | -84.4616
6.0337 | -84.4628
6.0344 | -84.464
6.0352 | -84.4659
6.036 | -84.468
6.0368 | -84.4665
6.0375 | -84.4666
6.0383 | -84.4678
6.0391 | -84.4684
6.0399 | -84.4691
6.0406 | -84.469
6.0414 | -84.4687
6.0422 | -84.4696
6.043 | -84.4703
6.0437 | -84.4714
6.0445 | -84.472
6.0453 | -84.4728
6.0461 | -84.4739
6.0468 | -84.4744
6.0476 | -84.4743
6.0484 | -84.4756
6.0491 | -84.4764
6.0499 | -84.4746
6.0507 | -84.4757
6.0515 | -84.4768
6.0522 | -84.4794
6.053 | -84.4834
6.0538 | -84.4849
6.0546 | -84.4856
6.0553 | -84.486
6.0561 | -84.4879
6.0569 | -84.4906
6.0577 | -84.4932
6.0584 | -84.4942
6.0592 | -84.496
6.06 | -84.4974
6.0608 | -84.4995
6.0615 | -84.5008
6.0623 | -84.5022
6.0631 | -84.5039
6.0639 | -84.5037
6.0646 | -84.503
6.0654 | -84.5039
6.0662 | -84.5053
6.067 | -84.5088
6.0677 | -84.5094
6.0685 | -84.5104
6.0693 | -84.512
6.0701 | -84.513
6.0708 | -84.5145
6.0716 | -84.5148
6.0724 | -84.5157
6.0732 | -84.5165
6.0739 | -84.5177
6.0747 | -84.5197
6.0755 | -84.5205
6.0763 | -84.5221
6.077 | -84.5228
6.0778 | -84.5242
6.0786 | -84.5249
6.0794 | -84.5274
6.0801 | -84.5296
6.0809 | -84.5317
6.0817 | -84.5337
6.0825 | -84.5367
6.0832 | -84.5381
6.084 | -84.5386
6.0848 | -84.5399
6.0856 | -84.5427
6.0863 | -84.5457
6.0871 | -84.5475
6.0879 | -84.549
6.0887 | -84.5509
6.0894 | -84.5526
6.0902 | -84.5528
6.091 | -84.5545
6.0918 | -84.5578
6.0925 | -84.5606
6.0933 | -84.5619
6.0941 | -84.5641
6.0949 | -84.5666
6.0956 | -84.57
6.0964 | -84.5753
6.0972 | -84.577
6.098 | -84.5791
6.0987 | -84.5826
6.0995 | -84.5828
6.1003 | -84.5827
6.1011 | -84.5862
6.1018 | -84.59
6.1026 | -84.5929
6.1034 | -84.5954
6.1042 | -84.5957
6.1049 | -84.595
6.1057 | -84.5971
6.1065 | -84.5995
6.1073 | -84.6025
6.108 | -84.6071
6.1088 | -84.6118
6.1096 | -84.6141
6.1104 | -84.616
6.1111 | -84.6183
6.1119 | -84.6207
6.1127 | -84.6237
6.1135 | -84.6269
6.1142 | -84.6291
6.115 | -84.6312
6.1158 | -84.6327
6.1166 | -84.6363
6.1173 | -84.641
6.1181 | -84.6437
6.1189 | -84.6454
6.1197 | -84.6483
6.1204 | -84.6513
6.1212 | -84.6544
6.122 | -84.6577
6.1227 | -84.6584
6.1235 | -84.661
6.1243 | -84.6656
6.1251 | -84.668
6.1258 | -84.6704
6.1266 | -84.6761
6.1274 | -84.6829
6.1282 | -84.6871
6.1289 | -84.689
6.1297 | -84.6915
6.1305 | -84.6952
6.1313 | -84.6997
6.132 | -84.7018
6.1328 | -84.7075
6.1336 | -84.7138
6.1344 | -84.716
6.1351 | -84.7164
6.1359 | -84.7171
6.1367 | -84.7203
6.1375 | -84.7229
6.1382 | -84.7255
6.139 | -84.7288
6.1398 | -84.7324
6.1406 | -84.7352
6.1413 | -84.7386
6.1421 | -84.7412
6.1429 | -84.7449
6.1437 | -84.7471
6.1444 | -84.7479
6.1452 | -84.7519
6.146 | -84.755
6.1468 | -84.7595
6.1475 | -84.7611
6.1483 | -84.7633
6.1491 | -84.7674
6.1499 | -84.7712
6.1506 | -84.7742
6.1514 | -84.7783
6.1522 | -84.7813
6.153 | -84.7837
6.1537 | -84.7869
6.1545 | -84.7896
6.1553 | -84.7934
6.1561 | -84.8
6.1568 | -84.8047
6.1576 | -84.8077
6.1584 | -84.81
6.1592 | -84.8141
6.1599 | -84.8189
6.1607 | -84.8231
6.1615 | -84.8267
6.1623 | -84.8305
6.163 | -84.8338
6.1638 | -84.8356
6.1646 | -84.8393
6.1654 | -84.8428
6.1661 | -84.8466
6.1669 | -84.8508
6.1677 | -84.8548
6.1685 | -84.8583
6.1692 | -84.8598
6.17 | -84.8611
6.1708 | -84.8647
6.1716 | -84.868
6.1723 | -84.8716
6.1731 | -84.8754
6.1739 | -84.8783
6.1747 | -84.881
6.1754 | -84.883
6.1762 | -84.8853
6.177 | -84.889
6.1778 | -84.8915
6.1785 | -84.8935
6.1793 | -84.8952
6.1801 | -84.9
6.1809 | -84.9053
6.1816 | -84.9108
6.1824 | -84.9163
6.1832 | -84.9192
6.184 | -84.9225
6.1847 | -84.9271
6.1855 | -84.9317
6.1863 | -84.9346
6.1871 | -84.9371
6.1878 | -84.943
6.1886 | -84.948
6.1894 | -84.9517
6.1902 | -84.9551
6.1909 | -84.9586
6.1917 | -84.9635
6.1925 | -84.966
6.1932 | -84.9676
6.194 | -84.97
6.1948 | -84.9745
6.1956 | -84.9754
6.1963 | -84.9785
6.1971 | -84.9811
6.1979 | -84.9842
6.1987 | -84.9868
6.1994 | -84.9905
6.2002 | -84.9933
6.201 | -84.9962
6.2018 | -84.9989
6.2025 | -85.0017
6.2033 | -85.0072
6.2041 | -85.0148
6.2049 | -85.0183
6.2056 | -85.0222
6.2064 | -85.0255
6.2072 | -85.0292
6.208 | -85.0324
6.2087 | -85.0354
6.2095 | -85.039
6.2103 | -85.0444
6.2111 | -85.0483
6.2118 | -85.0513
6.2126 | -85.0547
6.2134 | -85.0589
6.2142 | -85.0629
6.2149 | -85.0664
6.2157 | -85.0702
6.2165 | -85.0736
6.2173 | -85.0769
6.218 | -85.0811
6.2188 | -85.0845
6.2196 | -85.0881
6.2204 | -85.0911
6.2211 | -85.0952
6.2219 | -85.0992
6.2227 | -85.1022
6.2235 | -85.1051
6.2242 | -85.1079
6.225 | -85.1096
6.2258 | -85.1122
6.2266 | -85.1176
6.2273 | -85.1231
6.2281 | -85.1266
6.2289 | -85.1304
6.2297 | -85.1346
6.2304 | -85.1371
6.2312 | -85.143
6.232 | -85.1472
6.2328 | -85.1501
6.2335 | -85.1533
6.2343 | -85.1574
6.2351 | -85.1617
6.2359 | -85.1669
6.2366 | -85.1705
6.2374 | -85.1728
6.2382 | -85.1765
6.239 | -85.1814
6.2397 | -85.1848
6.2405 | -85.1899
6.2413 | -85.1933
6.2421 | -85.1958
6.2428 | -85.1998
6.2436 | -85.2044
6.2444 | -85.2085
6.2452 | -85.211
6.2459 | -85.2124
6.2467 | -85.2146
6.2475 | -85.2173
6.2483 | -85.2241
6.249 | -85.2304
6.2498 | -85.2342
6.2506 | -85.2353
6.2514 | -85.2374
6.2521 | -85.2412
6.2529 | -85.2463
6.2537 | -85.2512
6.2545 | -85.2552
6.2552 | -85.2596
6.256 | -85.2648
6.2568 | -85.2698
6.2576 | -85.2738
6.2583 | -85.2765
6.2591 | -85.2804
6.2599 | -85.2843
6.2607 | -85.288
6.2614 | -85.2911
6.2622 | -85.296
6.263 | -85.3026
6.2638 | -85.3086
6.2645 | -85.3133
6.2653 | -85.3172
6.2661 | -85.321
6.2669 | -85.3233
6.2676 | -85.3251
6.2684 | -85.3268
6.2692 | -85.3294
6.2699 | -85.3321
6.2707 | -85.3355
6.2715 | -85.3415
6.2723 | -85.3462
6.273 | -85.3485
6.2738 | -85.3509
6.2746 | -85.3542
6.2754 | -85.3585
6.2761 | -85.3624
6.2769 | -85.3661
6.2777 | -85.3709
6.2785 | -85.3731
6.2792 | -85.3747
6.28 | -85.3753
6.2808 | -85.3783
6.2816 | -85.381
6.2823 | -85.3835
6.2831 | -85.3875
6.2839 | -85.391
6.2847 | -85.3948
6.2854 | -85.3998
6.2862 | -85.4052
6.287 | -85.4079
6.2878 | -85.4107
6.2885 | -85.4167
6.2893 | -85.4238
6.2901 | -85.4307
6.2909 | -85.4345
6.2916 | -85.4383
6.2924 | -85.4425
6.2932 | -85.4473
6.294 | -85.4519
6.2947 | -85.4575
6.2955 | -85.4618
6.2963 | -85.4648
6.2971 | -85.4684
6.2978 | -85.4719
6.2986 | -85.4746
6.2994 | -85.4783
6.3002 | -85.482
6.3009 | -85.486
6.3017 | -85.4901
6.3025 | -85.4922
6.3033 | -85.4918
6.304 | -85.4937
6.3048 | -85.4976
6.3056 | -85.5001
6.3064 | -85.5025
6.3071 | -85.506
6.3079 | -85.5091
6.3087 | -85.5131
6.3095 | -85.5211
6.3102 | -85.5305
6.311 | -85.5389
6.3118 | -85.543
6.3126 | -85.5473
6.3133 | -85.5537
6.3141 | -85.5575
6.3149 | -85.5607
6.3157 | -85.5638
6.3164 | -85.5667
6.3172 | -85.5716
6.318 | -85.5778
6.3188 | -85.583
6.3195 | -85.5856
6.3203 | -85.5885
6.3211 | -85.5923
6.3219 | -85.5953
6.3226 | -85.5983
6.3234 | -85.601
6.3242 | -85.604
6.325 | -85.6075
6.3257 | -85.6117
6.3265 | -85.6155
6.3273 | -85.615
6.3281 | -85.6163
6.3288 | -85.6194
6.3296 | -85.6243
6.3304 | -85.6265
6.3312 | -85.629
6.3319 | -85.6304
6.3327 | -85.6324
6.3335 | -85.6357
6.3343 | -85.638
6.335 | -85.6412
6.3358 | -85.6444
6.3366 | -85.6454
6.3374 | -85.6463
6.3381 | -85.6473
6.3389 | -85.6511
6.3397 | -85.6547
6.3405 | -85.6572
6.3412 | -85.6594
6.342 | -85.6621
6.3428 | -85.6659
6.3435 | -85.6696
6.3443 | -85.6718
6.3451 | -85.6729
6.3459 | -85.6746
6.3466 | -85.6776
6.3474 | -85.6804
6.3482 | -85.6838
6.349 | -85.6859
6.3497 | -85.6903
6.3505 | -85.694
6.3513 | -85.6976
6.3521 | -85.7017
6.3528 | -85.707
6.3536 | -85.7092
6.3544 | -85.7113
6.3552 | -85.7151
6.3559 | -85.7175
6.3567 | -85.7198
6.3575 | -85.7234
6.3583 | -85.7262
6.359 | -85.7285
6.3598 | -85.7306
6.3606 | -85.7321
6.3614 | -85.7346
6.3621 | -85.7357
6.3629 | -85.7342
6.3637 | -85.7338
6.3645 | -85.7356
6.3652 | -85.7386
6.366 | -85.7415
6.3668 | -85.7436
6.3676 | -85.7447
6.3683 | -85.7461
6.3691 | -85.7483
6.3699 | -85.7498
6.3707 | -85.752
6.3714 | -85.7539
6.3722 | -85.7574
6.373 | -85.7615
6.3738 | -85.7626
6.3745 | -85.7644
6.3753 | -85.7682
6.3761 | -85.7714
6.3769 | -85.7736
6.3776 | -85.776
6.3784 | -85.7783
6.3792 | -85.7808
6.38 | -85.7829
6.3807 | -85.7845
6.3815 | -85.7852
6.3823 | -85.7879
6.3831 | -85.7913
6.3838 | -85.792
6.3846 | -85.7908
6.3854 | -85.7871
6.3862 | -85.7863
6.3869 | -85.7878
6.3877 | -85.7902
6.3885 | -85.7944
6.3893 | -85.7952
6.39 | -85.7954
6.3908 | -85.7963
6.3916 | -85.7969
6.3924 | -85.7937
6.3931 | -85.794
6.3939 | -85.7962
6.3947 | -85.7999
6.3955 | -85.8013
6.3962 | -85.8019
6.397 | -85.8041
6.3978 | -85.8068
6.3986 | -85.808
6.3993 | -85.808
6.4001 | -85.8056
6.4009 | -85.8038
6.4017 | -85.8049
6.4024 | -85.8057
6.4032 | -85.8061
6.404 | -85.8061
6.4048 | -85.8078
6.4055 | -85.8126
6.4063 | -85.8134
6.4071 | -85.8142
6.4079 | -85.814
6.4086 | -85.8145
6.4094 | -85.8135
6.4102 | -85.8113
6.411 | -85.8096
6.4117 | -85.8091
6.4125 | -85.81
6.4133 | -85.8115
6.4141 | -85.8135
6.4148 | -85.8129
6.4156 | -85.8123
6.4164 | -85.8111
6.4172 | -85.8081
6.4179 | -85.8067
6.4187 | -85.8065
6.4195 | -85.8061
6.4202 | -85.8054
6.421 | -85.8067
6.4218 | -85.8064
6.4226 | -85.8059
6.4233 | -85.8057
6.4241 | -85.8059
6.4249 | -85.8052
6.4257 | -85.8044
6.4264 | -85.8042
6.4272 | -85.8058
6.428 | -85.8059
6.4288 | -85.8052
6.4295 | -85.8063
6.4303 | -85.8091
6.4311 | -85.8105
6.4319 | -85.8108
6.4326 | -85.8102
6.4334 | -85.8074
6.4342 | -85.8061
6.435 | -85.8051
6.4357 | -85.804
6.4365 | -85.8035
6.4373 | -85.8029
6.4381 | -85.8038
6.4388 | -85.8038
6.4396 | -85.8031
6.4404 | -85.8014
6.4412 | -85.799
6.4419 | -85.798
6.4427 | -85.7965
6.4435 | -85.7945
6.4443 | -85.7934
6.445 | -85.792
6.4458 | -85.7909
6.4466 | -85.79
6.4474 | -85.7891
6.4481 | -85.7873
6.4489 | -85.7863
6.4497 | -85.7857
6.4505 | -85.7849
6.4512 | -85.7841
6.452 | -85.7842
6.4528 | -85.7815
6.4536 | -85.7802
6.4543 | -85.7811
6.4551 | -85.7835
6.4559 | -85.7823
6.4567 | -85.7821
6.4574 | -85.7813
6.4582 | -85.7799
6.459 | -85.7784
6.4598 | -85.7776
6.4605 | -85.7767
6.4613 | -85.7756
6.4621 | -85.7748
6.4629 | -85.7756
6.4636 | -85.7746
6.4644 | -85.7733
6.4652 | -85.7731
6.466 | -85.773
6.4667 | -85.7722
6.4675 | -85.7695
6.4683 | -85.7685
6.4691 | -85.7675
6.4698 | -85.7651
6.4706 | -85.7597
6.4714 | -85.7585
6.4722 | -85.7573
6.4729 | -85.7557
6.4737 | -85.7542
6.4745 | -85.7531
6.4753 | -85.7517
6.476 | -85.7527
6.4768 | -85.755
6.4776 | -85.7548
6.4784 | -85.7537
6.4791 | -85.7519
6.4799 | -85.7499
6.4807 | -85.7467
6.4815 | -85.7444
6.4822 | -85.743
6.483 | -85.7422
6.4838 | -85.7416
6.4846 | -85.7399
6.4853 | -85.7395
6.4861 | -85.7414
6.4869 | -85.7394
6.4877 | -85.7364
6.4884 | -85.7334
6.4892 | -85.7316
6.49 | -85.7293
6.4908 | -85.7283
6.4915 | -85.7279
6.4923 | -85.7254
6.4931 | -85.7213
6.4938 | -85.7185
6.4946 | -85.716
6.4954 | -85.7144
6.4962 | -85.7125
6.4969 | -85.7113
6.4977 | -85.7091
6.4985 | -85.7068
6.4993 | -85.7048
6.5 | -85.7036
6.5008 | -85.7041
6.5016 | -85.7045
6.5024 | -85.7036
6.5031 | -85.7013
6.5039 | -85.6983
6.5047 | -85.698
6.5055 | -85.6966
6.5062 | -85.6953
6.507 | -85.6938
6.5078 | -85.6921
6.5086 | -85.6899
6.5093 | -85.6896
6.5101 | -85.6895
6.5109 | -85.6859
6.5117 | -85.6843
6.5124 | -85.6822
6.5132 | -85.6813
6.514 | -85.68
6.5148 | -85.6795
6.5155 | -85.6784
6.5163 | -85.6782
6.5171 | -85.6782
6.5179 | -85.6759
6.5186 | -85.6736
6.5194 | -85.6721
6.5202 | -85.671
6.521 | -85.67
6.5217 | -85.6694
6.5225 | -85.6699
6.5233 | -85.6695
6.5241 | -85.6676
6.5248 | -85.6659
6.5256 | -85.6641
6.5264 | -85.6608
6.5272 | -85.6601
6.5279 | -85.6585
6.5287 | -85.6569
6.5295 | -85.6557
6.5303 | -85.6541
6.531 | -85.6528
6.5318 | -85.6511
6.5326 | -85.6504
6.5334 | -85.6493
6.5341 | -85.6478
6.5349 | -85.6466
6.5357 | -85.6454
6.5365 | -85.6449
6.5372 | -85.6444
6.538 | -85.6439
6.5388 | -85.6459
6.5396 | -85.6477
6.5403 | -85.6473
6.5411 | -85.6461
6.5419 | -85.6458
6.5427 | -85.6467
6.5434 | -85.6456
6.5442 | -85.6458
6.545 | -85.6477
6.5458 | -85.6473
6.5465 | -85.6463
6.5473 | -85.6456
6.5481 | -85.6452
6.5489 | -85.6473
6.5496 | -85.6474
6.5504 | -85.6492
6.5512 | -85.65
6.552 | -85.6487
6.5527 | -85.6469
6.5535 | -85.646
6.5543 | -85.6442
6.5551 | -85.6411
6.5558 | -85.6417
6.5566 | -85.6419
6.5574 | -85.6402
6.5582 | -85.6387
6.5589 | -85.6389
6.5597 | -85.6395
6.5605 | -85.6391
6.5613 | -85.6375
6.562 | -85.6368
6.5628 | -85.6365
6.5636 | -85.639
6.5644 | -85.6405
6.5651 | -85.6401
6.5659 | -85.6393
6.5667 | -85.6397
6.5674 | -85.6409
6.5682 | -85.6409
6.569 | -85.6393
6.5698 | -85.6371
6.5705 | -85.6351
6.5713 | -85.6368
6.5721 | -85.6378
6.5729 | -85.6371
6.5736 | -85.6362
6.5744 | -85.6365
6.5752 | -85.6366
6.576 | -85.6366
6.5767 | -85.6364
6.5775 | -85.6368
6.5783 | -85.6381
6.5791 | -85.6391
6.5798 | -85.6384
6.5806 | -85.6386
6.5814 | -85.6395
6.5822 | -85.6385
6.5829 | -85.6386
6.5837 | -85.639
6.5845 | -85.6399
6.5853 | -85.6406
6.586 | -85.6424
6.5868 | -85.6448
6.5876 | -85.6455
6.5884 | -85.6462
6.5891 | -85.6466
6.5899 | -85.6438
6.5907 | -85.6432
6.5915 | -85.6452
6.5922 | -85.6458
6.593 | -85.6481
6.5938 | -85.6515
6.5946 | -85.6532
6.5953 | -85.6537
6.5961 | -85.6542
6.5969 | -85.6541
6.5977 | -85.6544
6.5984 | -85.6555
6.5992 | -85.6566
6.6 | -85.6573
6.6008 | -85.6572
6.6015 | -85.658
6.6023 | -85.6593
6.6031 | -85.6606
6.6039 | -85.662
6.6046 | -85.6625
6.6054 | -85.6639
6.6062 | -85.6656
6.607 | -85.6688
6.6077 | -85.6722
6.6085 | -85.6748
6.6093 | -85.6769
6.6101 | -85.6788
6.6108 | -85.6799
6.6116 | -85.6811
6.6124 | -85.6824
6.6132 | -85.6853
6.6139 | -85.6897
6.6147 | -85.6931
6.6155 | -85.6937
6.6163 | -85.6932
6.617 | -85.6947
6.6178 | -85.6965
6.6186 | -85.6986
6.6194 | -85.6983
6.6201 | -85.6969
6.6209 | -85.6992
6.6217 | -85.7023
6.6225 | -85.7051
6.6232 | -85.7077
6.624 | -85.7085
6.6248 | -85.711
6.6256 | -85.7128
6.6263 | -85.7141
6.6271 | -85.716
6.6279 | -85.718
6.6287 | -85.7205
6.6294 | -85.7219
6.6302 | -85.7242
6.631 | -85.727
6.6318 | -85.73
6.6325 | -85.734
6.6333 | -85.7364
6.6341 | -85.7386
6.6349 | -85.7413
6.6356 | -85.7439
6.6364 | -85.7462
6.6372 | -85.7484
6.638 | -85.7503
6.6387 | -85.7551
6.6395 | -85.7582
6.6403 | -85.7626
6.641 | -85.7643
6.6418 | -85.7668
6.6426 | -85.7717
6.6434 | -85.7761
6.6441 | -85.7796
6.6449 | -85.7822
6.6457 | -85.7839
6.6465 | -85.7859
6.6472 | -85.7889
6.648 | -85.7915
6.6488 | -85.7947
6.6496 | -85.7977
6.6503 | -85.8005
6.6511 | -85.8028
6.6519 | -85.8026
6.6527 | -85.8024
6.6534 | -85.8049
6.6542 | -85.8073
6.655 | -85.8086
6.6558 | -85.8082
6.6565 | -85.8069
6.6573 | -85.8081
6.6581 | -85.8112
6.6589 | -85.814
6.6596 | -85.8152
6.6604 | -85.8179
6.6612 | -85.8209
6.662 | -85.8237
6.6627 | -85.8268
6.6635 | -85.8287
6.6643 | -85.8313
6.6651 | -85.8346
6.6658 | -85.8372
6.6666 | -85.8394
6.6674 | -85.8427
6.6682 | -85.8453
6.6689 | -85.8491
6.6697 | -85.8497
6.6705 | -85.8516
6.6713 | -85.8547
6.672 | -85.8574
6.6728 | -85.8603
6.6736 | -85.8654
6.6744 | -85.8701
6.6751 | -85.873
6.6759 | -85.8748
6.6767 | -85.8774
6.6775 | -85.8803
6.6782 | -85.883
6.679 | -85.8866
6.6798 | -85.8896
6.6806 | -85.8924
6.6813 | -85.8958
6.6821 | -85.8991
6.6829 | -85.9018
6.6837 | -85.9038
6.6844 | -85.9031
6.6852 | -85.9038
6.686 | -85.9067
6.6868 | -85.9086
6.6875 | -85.911
6.6883 | -85.9125
6.6891 | -85.9142
6.6899 | -85.917
6.6906 | -85.9197
6.6914 | -85.9227
6.6922 | -85.9274
6.693 | -85.9325
6.6937 | -85.9355
6.6945 | -85.938
6.6953 | -85.9411
6.6961 | -85.9445
6.6968 | -85.9476
6.6976 | -85.9496
6.6984 | -85.9521
6.6992 | -85.9556
6.6999 | -85.9578
6.7007 | -85.961
6.7015 | -85.9641
6.7023 | -85.966
6.703 | -85.968
6.7038 | -85.9697
6.7046 | -85.9711
6.7054 | -85.9722
6.7061 | -85.9745
6.7069 | -85.9762
6.7077 | -85.9772
6.7085 | -85.9788
6.7092 | -85.9817
6.71 | -85.9857
6.7108 | -85.9869
6.7115 | -85.9881
6.7123 | -85.9905
6.7131 | -85.9938
6.7139 | -85.9962
6.7146 | -85.9986
6.7154 | -86.0005
6.7162 | -86.002
6.717 | -86.0024
6.7177 | -86.0035
6.7185 | -86.0071
6.7193 | -86.0111
6.7201 | -86.0152
6.7208 | -86.0185
6.7216 | -86.0211
6.7224 | -86.0219
6.7232 | -86.0221
6.7239 | -86.0228
6.7247 | -86.0254
6.7255 | -86.0294
6.7263 | -86.0311
6.727 | -86.0331
6.7278 | -86.0348
6.7286 | -86.0349
6.7294 | -86.0352
6.7301 | -86.0369
6.7309 | -86.0397
6.7317 | -86.0434
6.7325 | -86.0485
6.7332 | -86.0516
6.734 | -86.0529
6.7348 | -86.0557
6.7356 | -86.0582
6.7363 | -86.0597
6.7371 | -86.0612
6.7379 | -86.0627
6.7387 | -86.0637
6.7394 | -86.0651
6.7402 | -86.0678
6.741 | -86.0707
6.7418 | -86.0726
6.7425 | -86.0745
6.7433 | -86.0756
6.7441 | -86.077
6.7449 | -86.0795
6.7456 | -86.0821
6.7464 | -86.0832
6.7472 | -86.0845
6.748 | -86.0865
6.7487 | -86.0896
6.7495 | -86.0924
6.7503 | -86.0918
6.7511 | -86.0933
6.7518 | -86.0939
6.7526 | -86.0948
6.7534 | -86.0965
6.7542 | -86.0971
6.7549 | -86.0969
6.7557 | -86.1
6.7565 | -86.103
6.7573 | -86.1067
6.758 | -86.1102
6.7588 | -86.1125
6.7596 | -86.1137
6.7604 | -86.1162
6.7611 | -86.1184
6.7619 | -86.1195
6.7627 | -86.1202
6.7635 | -86.1209
6.7642 | -86.122
6.765 | -86.1227
6.7658 | -86.1236
6.7666 | -86.1231
6.7673 | -86.122
6.7681 | -86.1227
6.7689 | -86.1238
6.7697 | -86.1237
6.7704 | -86.1225
6.7712 | -86.1232
6.772 | -86.125
6.7728 | -86.124
6.7735 | -86.1253
6.7743 | -86.1246
6.7751 | -86.1259
6.7759 | -86.1273
6.7766 | -86.1282
6.7774 | -86.1293
6.7782 | -86.1302
6.779 | -86.1303
6.7797 | -86.1293
6.7805 | -86.1297
6.7813 | -86.1298
6.7821 | -86.1305
6.7828 | -86.1306
6.7836 | -86.1344
6.7844 | -86.137
6.7851 | -86.1375
6.7859 | -86.1379
6.7867 | -86.1391
6.7875 | -86.1398
6.7882 | -86.1416
6.789 | -86.1423
6.7898 | -86.1413
6.7906 | -86.1414
6.7913 | -86.1406
6.7921 | -86.1408
6.7929 | -86.1429
6.7937 | -86.1443
6.7944 | -86.1464
6.7952 | -86.147
6.796 | -86.1469
6.7968 | -86.1469
6.7975 | -86.1484
6.7983 | -86.149
6.7991 | -86.1495
6.7999 | -86.1503
6.8006 | -86.1507
6.8014 | -86.1517
6.8022 | -86.1544
6.803 | -86.1564
6.8037 | -86.1567
6.8045 | -86.1563
6.8053 | -86.1567
6.8061 | -86.1568
6.8068 | -86.1574
6.8076 | -86.1572
6.8084 | -86.1567
6.8092 | -86.1571
6.8099 | -86.1582
6.8107 | -86.1585
6.8115 | -86.1602
6.8123 | -86.1615
6.813 | -86.1604
6.8138 | -86.1582
6.8146 | -86.1578
6.8154 | -86.1582
6.8161 | -86.1587
6.8169 | -86.1587
6.8177 | -86.1592
6.8185 | -86.1592
6.8192 | -86.1606
6.82 | -86.1604
6.8208 | -86.1593
6.8216 | -86.1611
6.8223 | -86.1639
6.8231 | -86.1659
6.8239 | -86.1659
6.8247 | -86.1665
6.8254 | -86.1673
6.8262 | -86.1693
6.827 | -86.17
6.8278 | -86.1699
6.8285 | -86.1708
6.8293 | -86.1677
6.8301 | -86.1662
6.8309 | -86.1668
6.8316 | -86.1678
6.8324 | -86.1678
6.8332 | -86.1688
6.834 | -86.169
6.8347 | -86.1694
6.8355 | -86.1695
6.8363 | -86.1701
6.8371 | -86.1701
6.8378 | -86.1701
6.8386 | -86.1708
6.8394 | -86.171
6.8402 | -86.1718
6.8409 | -86.1727
6.8417 | -86.1742
6.8425 | -86.1756
6.8433 | -86.1753
6.844 | -86.1762
6.8448 | -86.1788
6.8456 | -86.1841
6.8464 | -86.1869
6.8471 | -86.1873
6.8479 | -86.1878
6.8487 | -86.1883
6.8495 | -86.189
6.8502 | -86.1905
6.851 | -86.1918
6.8518 | -86.1923
6.8526 | -86.1953
6.8533 | -86.1974
6.8541 | -86.1992
6.8549 | -86.1992
6.8557 | -86.1997
6.8564 | -86.2008
6.8572 | -86.2
6.858 | -86.1989
6.8587 | -86.1982
6.8595 | -86.1967
6.8603 | -86.1973
6.8611 | -86.2006
6.8618 | -86.2002
6.8626 | -86.2018
6.8634 | -86.1994
6.8642 | -86.1984
6.8649 | -86.1991
6.8657 | -86.2001
6.8665 | -86.2009
6.8673 | -86.2004
6.868 | -86.2033
6.8688 | -86.2062
6.8696 | -86.207
6.8704 | -86.2069
6.8711 | -86.2077
6.8719 | -86.2104
6.8727 | -86.2137
6.8735 | -86.2156
6.8742 | -86.216
6.875 | -86.2168
6.8758 | -86.2182
6.8766 | -86.2192
6.8773 | -86.2199
6.8781 | -86.2212
6.8789 | -86.2228
6.8797 | -86.2265
6.8804 | -86.2287
6.8812 | -86.2286
6.882 | -86.2283
6.8828 | -86.2302
6.8835 | -86.2307
6.8843 | -86.234
6.8851 | -86.2358
6.8859 | -86.2362
6.8866 | -86.2355
6.8874 | -86.2352
6.8882 | -86.2362
6.889 | -86.2391
6.8897 | -86.2408
6.8905 | -86.2409
6.8913 | -86.2402
6.8921 | -86.2412
6.8928 | -86.2415
6.8936 | -86.2415
6.8944 | -86.2425
6.8952 | -86.2432
6.8959 | -86.243
6.8967 | -86.2462
6.8975 | -86.2479
6.8983 | -86.2489
6.899 | -86.2502
6.8998 | -86.2517
6.9006 | -86.253
6.9014 | -86.2541
6.9021 | -86.2558
6.9029 | -86.258
6.9037 | -86.2595
6.9045 | -86.2604
6.9052 | -86.2621
6.906 | -86.2635
6.9068 | -86.2675
6.9076 | -86.2696
6.9083 | -86.2711
6.9091 | -86.2723
6.9099 | -86.2743
6.9107 | -86.2762
6.9114 | -86.2783
6.9122 | -86.2805
6.913 | -86.2818
6.9138 | -86.2842
6.9145 | -86.2854
6.9153 | -86.286
6.9161 | -86.2878
6.9169 | -86.2907
6.9176 | -86.2945
6.9184 | -86.2966
6.9192 | -86.2986
6.92 | -86.3002
6.9207 | -86.3019
6.9215 | -86.3025
6.9223 | -86.3026
6.9231 | -86.3038
6.9238 | -86.3058
6.9246 | -86.3082
6.9254 | -86.3097
6.9262 | -86.3114
6.9269 | -86.3141
6.9277 | -86.317
6.9285 | -86.3195
6.9293 | -86.3233
6.93 | -86.3254
6.9308 | -86.3276
6.9316 | -86.3302
6.9323 | -86.332
6.9331 | -86.3335
6.9339 | -86.3363
6.9347 | -86.3378
6.9354 | -86.3392
6.9362 | -86.3433
6.937 | -86.3509
6.9378 | -86.3541
6.9385 | -86.3582
6.9393 | -86.3611
6.9401 | -86.3633
6.9409 | -86.3648
6.9416 | -86.3669
6.9424 | -86.3699
6.9432 | -86.3728
6.944 | -86.3745
6.9447 | -86.3755
6.9455 | -86.3775
6.9463 | -86.3804
6.9471 | -86.3846
6.9478 | -86.3882
6.9486 | -86.3907
6.9494 | -86.3943
6.9502 | -86.3982
6.9509 | -86.4009
6.9517 | -86.4012
6.9525 | -86.4039
6.9533 | -86.4063
6.954 | -86.408
6.9548 | -86.411
6.9556 | -86.4147
6.9564 | -86.4171
6.9571 | -86.4183
6.9579 | -86.4216
6.9587 | -86.4242
6.9595 | -86.4267
6.9602 | -86.4302
6.961 | -86.435
6.9618 | -86.4388
6.9626 | -86.4425
6.9633 | -86.4466
6.9641 | -86.4492
6.9649 | -86.4507
6.9657 | -86.4533
6.9664 | -86.4551
6.9672 | -86.4571
6.968 | -86.4578
6.9688 | -86.457
6.9695 | -86.4596
6.9703 | -86.4625
6.9711 | -86.4661
6.9719 | -86.4702
6.9726 | -86.4729
6.9734 | -86.4745
6.9742 | -86.4771
6.975 | -86.4806
6.9757 | -86.4796
6.9765 | -86.4799
6.9773 | -86.4832
6.9781 | -86.4856
6.9788 | -86.488
6.9796 | -86.4867
6.9804 | -86.4873
6.9812 | -86.4874
6.9819 | -86.4901
6.9827 | -86.4933
6.9835 | -86.4954
6.9843 | -86.4976
6.985 | -86.4991
6.9858 | -86.5012
6.9866 | -86.5037
6.9874 | -86.5071
6.9881 | -86.5104
6.9889 | -86.5131
6.9897 | -86.5149
6.9905 | -86.516
6.9912 | -86.5172
6.992 | -86.5189
6.9928 | -86.5213
6.9936 | -86.5234
6.9943 | -86.5244
6.9951 | -86.5253
6.9959 | -86.526
6.9967 | -86.5314
6.9974 | -86.5352
6.9982 | -86.5372
6.999 | -86.5402
6.9998 | -86.5437
7.0005 | -86.548
7.0013 | -86.551
7.0021 | -86.552
7.0029 | -86.5557
7.0036 | -86.559
7.0044 | -86.5616
7.0052 | -86.5643
7.006 | -86.5661
7.0067 | -86.568
7.0075 | -86.5703
7.0083 | -86.5729
7.009 | -86.575
7.0098 | -86.5774
7.0106 | -86.5801
7.0114 | -86.5815
7.0121 | -86.5837
7.0129 | -86.5856
7.0137 | -86.5876
7.0145 | -86.5903
7.0152 | -86.5943
7.016 | -86.598
7.0168 | -86.5995
7.0176 | -86.602
7.0183 | -86.6023
7.0191 | -86.6009
7.0199 | -86.6028
7.0207 | -86.6049
7.0214 | -86.607
7.0222 | -86.6092
7.023 | -86.6104
7.0238 | -86.6104
7.0245 | -86.6128
7.0253 | -86.6169
7.0261 | -86.6201
7.0269 | -86.6236
7.0276 | -86.6284
7.0284 | -86.6307
7.0292 | -86.6323
7.03 | -86.6358
7.0307 | -86.6381
7.0315 | -86.6396
7.0323 | -86.6413
7.0331 | -86.6413
7.0338 | -86.6427
7.0346 | -86.6446
7.0354 | -86.6459
7.0362 | -86.6471
7.0369 | -86.6482
7.0377 | -86.6498
7.0385 | -86.6507
7.0393 | -86.652
7.04 | -86.6525
7.0408 | -86.653
7.0416 | -86.6547
7.0424 | -86.658
7.0431 | -86.6627
7.0439 | -86.6671
7.0447 | -86.6698
7.0455 | -86.6714
7.0462 | -86.6719
7.047 | -86.6734
7.0478 | -86.6761
7.0486 | -86.6778
7.0493 | -86.6793
7.0501 | -86.6808
7.0509 | -86.6835
7.0517 | -86.687
7.0524 | -86.6893
7.0532 | -86.6904
7.054 | -86.6929
7.0548 | -86.6938
7.0555 | -86.6944
7.0563 | -86.6965
7.0571 | -86.6973
7.0579 | -86.6978
7.0586 | -86.6978
7.0594 | -86.6983
7.0602 | -86.6997
7.061 | -86.702
7.0617 | -86.7038
7.0625 | -86.7053
7.0633 | -86.7068
7.0641 | -86.7085
7.0648 | -86.7135
7.0656 | -86.7176
7.0664 | -86.7187
7.0672 | -86.7203
7.0679 | -86.7226
7.0687 | -86.7236
7.0695 | -86.7242
7.0703 | -86.7254
7.071 | -86.7251
7.0718 | -86.7228
7.0726 | -86.7245
7.0734 | -86.7251
7.0741 | -86.7257
7.0749 | -86.7264
7.0757 | -86.7272
7.0765 | -86.7296
7.0772 | -86.7304
7.078 | -86.7316
7.0788 | -86.7314
7.0796 | -86.7305
7.0803 | -86.7311
7.0811 | -86.7317
7.0819 | -86.7313
7.0826 | -86.7317
7.0834 | -86.7325
7.0842 | -86.7332
7.085 | -86.7329
7.0857 | -86.7313
7.0865 | -86.7315
7.0873 | -86.7328
7.0881 | -86.7335
7.0888 | -86.7342
7.0896 | -86.7337
7.0904 | -86.7338
7.0912 | -86.7344
7.0919 | -86.7369
7.0927 | -86.7419
7.0935 | -86.743
7.0943 | -86.7437
7.095 | -86.7443
7.0958 | -86.7442
7.0966 | -86.7472
7.0974 | -86.7517
7.0981 | -86.7519
7.0989 | -86.7513
7.0997 | -86.7502
7.1005 | -86.7495
7.1012 | -86.7511
7.102 | -86.7503
7.1028 | -86.7488
7.1036 | -86.749
7.1043 | -86.7488
7.1051 | -86.7476
7.1059 | -86.746
7.1067 | -86.746
7.1074 | -86.7466
7.1082 | -86.7468
7.109 | -86.7472
7.1098 | -86.7467
7.1105 | -86.7454
7.1113 | -86.745
7.1121 | -86.7453
7.1129 | -86.7451
7.1136 | -86.7447
7.1144 | -86.7442
7.1152 | -86.7438
7.116 | -86.7444
7.1167 | -86.744
7.1175 | -86.7431
7.1183 | -86.7434
7.1191 | -86.7422
7.1198 | -86.7439
7.1206 | -86.7448
7.1214 | -86.7446
7.1222 | -86.7442
7.1229 | -86.7425
7.1237 | -86.742
7.1245 | -86.7432
7.1253 | -86.7433
7.126 | -86.7431
7.1268 | -86.7418
7.1276 | -86.7411
7.1284 | -86.7413
7.1291 | -86.7409
7.1299 | -86.7416
7.1307 | -86.7413
7.1315 | -86.7449
7.1322 | -86.7442
7.133 | -86.7414
7.1338 | -86.7415
7.1346 | -86.7423
7.1353 | -86.7443
7.1361 | -86.7474
7.1369 | -86.7484
7.1377 | -86.7477
7.1384 | -86.7458
7.1392 | -86.7459
7.14 | -86.7452
7.1408 | -86.7449
7.1415 | -86.7452
7.1423 | -86.7449
7.1431 | -86.7427
7.1439 | -86.7404
7.1446 | -86.7385
7.1454 | -86.7378
7.1462 | -86.738
7.147 | -86.7373
7.1477 | -86.7364
7.1485 | -86.7348
7.1493 | -86.7337
7.1501 | -86.7354
7.1508 | -86.7343
7.1516 | -86.733
7.1524 | -86.733
7.1532 | -86.7333
7.1539 | -86.7328
7.1547 | -86.7313
7.1555 | -86.729
7.1562 | -86.7276
7.157 | -86.7288
7.1578 | -86.7296
7.1586 | -86.728
7.1593 | -86.7262
7.1601 | -86.7246
7.1609 | -86.7236
7.1617 | -86.7227
7.1624 | -86.7216
7.1632 | -86.7206
7.164 | -86.7201
7.1648 | -86.7201
7.1655 | -86.7201
7.1663 | -86.7236
7.1671 | -86.7249
7.1679 | -86.7242
7.1686 | -86.7252
7.1694 | -86.7261
7.1702 | -86.7257
7.171 | -86.7254
7.1717 | -86.7253
7.1725 | -86.7246
7.1733 | -86.7242
7.1741 | -86.7238
7.1748 | -86.7242
7.1756 | -86.7234
7.1764 | -86.7219
7.1772 | -86.7213
7.1779 | -86.7216
7.1787 | -86.7228
7.1795 | -86.7238
7.1803 | -86.7241
7.181 | -86.7239
7.1818 | -86.7219
7.1826 | -86.7201
7.1834 | -86.7214
7.1841 | -86.7224
7.1849 | -86.7221
7.1857 | -86.7214
7.1865 | -86.7215
7.1872 | -86.7209
7.188 | -86.7213
7.1888 | -86.7211
7.1896 | -86.7193
7.1903 | -86.718
7.1911 | -86.7186
7.1919 | -86.7181
7.1927 | -86.7173
7.1934 | -86.7168
7.1942 | -86.7174
7.195 | -86.717
7.1958 | -86.7154
7.1965 | -86.7146
7.1973 | -86.7115
7.1981 | -86.7067
7.1989 | -86.7039
7.1996 | -86.7019
7.2004 | -86.7017
7.2012 | -86.703
7.202 | -86.7038
7.2027 | -86.7025
7.2035 | -86.7014
7.2043 | -86.7014
7.2051 | -86.7008
7.2058 | -86.7006
7.2066 | -86.7015
7.2074 | -86.703
7.2082 | -86.7036
7.2089 | -86.7042
7.2097 | -86.7029
7.2105 | -86.7025
7.2113 | -86.702
7.212 | -86.7013
7.2128 | -86.7015
7.2136 | -86.7024
7.2144 | -86.7033
7.2151 | -86.7028
7.2159 | -86.7024
7.2167 | -86.6999
7.2175 | -86.699
7.2182 | -86.6986
7.219 | -86.6986
7.2198 | -86.6999
7.2206 | -86.7002
7.2213 | -86.7005
7.2221 | -86.7002
7.2229 | -86.6995
7.2237 | -86.6992
7.2244 | -86.6997
7.2252 | -86.6999
7.226 | -86.6985
7.2268 | -86.6977
7.2275 | -86.6968
7.2283 | -86.6958
7.2291 | -86.6963
7.2298 | -86.6964
7.2306 | -86.6954
7.2314 | -86.6947
7.2322 | -86.6952
7.2329 | -86.6953
7.2337 | -86.6953
7.2345 | -86.6952
7.2353 | -86.6945
7.236 | -86.6941
7.2368 | -86.6961
7.2376 | -86.696
7.2384 | -86.6959
7.2391 | -86.6946
7.2399 | -86.6923
7.2407 | -86.6926
7.2415 | -86.6948
7.2422 | -86.6964
7.243 | -86.6962
7.2438 | -86.6952
7.2446 | -86.6935
7.2453 | -86.6931
7.2461 | -86.6928
7.2469 | -86.6922
7.2477 | -86.6923
7.2484 | -86.693
7.2492 | -86.693
7.25 | -86.6931
7.2508 | -86.6921
7.2515 | -86.6904
7.2523 | -86.6903
7.2531 | -86.6896
7.2539 | -86.691
7.2546 | -86.693
7.2554 | -86.6929
7.2562 | -86.6928
7.257 | -86.6938
7.2577 | -86.696
7.2585 | -86.6962
7.2593 | -86.6947
7.2601 | -86.6933
7.2608 | -86.6925
7.2616 | -86.6944
7.2624 | -86.6966
7.2632 | -86.6963
7.2639 | -86.6974
7.2647 | -86.6982
7.2655 | -86.6982
7.2663 | -86.697
7.267 | -86.696
7.2678 | -86.6963
7.2686 | -86.6972
7.2694 | -86.6969
7.2701 | -86.6972
7.2709 | -86.6972
7.2717 | -86.6972
7.2725 | -86.6967
7.2732 | -86.6968
7.274 | -86.6972
7.2748 | -86.6971
7.2756 | -86.6973
7.2763 | -86.6977
7.2771 | -86.6966
7.2779 | -86.6949
7.2787 | -86.6962
7.2794 | -86.6971
7.2802 | -86.6977
7.281 | -86.6966
7.2818 | -86.6962
7.2825 | -86.6977
7.2833 | -86.6977
7.2841 | -86.6984
7.2849 | -86.6987
7.2856 | -86.6996
7.2864 | -86.7009
7.2872 | -86.7013
7.288 | -86.7018
7.2887 | -86.7022
7.2895 | -86.7025
7.2903 | -86.7028
7.2911 | -86.7031
7.2918 | -86.7034
7.2926 | -86.7039
7.2934 | -86.7043
7.2942 | -86.7056
7.2949 | -86.7046
7.2957 | -86.7035
7.2965 | -86.7033
7.2973 | -86.7039
7.298 | -86.7047
7.2988 | -86.7066
7.2996 | -86.7072
7.3003 | -86.7078
7.3011 | -86.708
7.3019 | -86.7089
7.3027 | -86.7091
7.3034 | -86.7111
7.3042 | -86.7147
7.305 | -86.7164
7.3058 | -86.7171
7.3065 | -86.718
7.3073 | -86.7195
7.3081 | -86.7209
7.3089 | -86.7231
7.3096 | -86.7232
7.3104 | -86.7242
7.3112 | -86.7252
7.312 | -86.7265
7.3127 | -86.7301
7.3135 | -86.7328
7.3143 | -86.7342
7.3151 | -86.7354
7.3158 | -86.7343
7.3166 | -86.7312
7.3174 | -86.7307
7.3182 | -86.7315
7.3189 | -86.7331
7.3197 | -86.7343
7.3205 | -86.7349
7.3213 | -86.7387
7.322 | -86.7408
7.3228 | -86.7407
7.3236 | -86.7418
7.3244 | -86.742
7.3251 | -86.7423
7.3259 | -86.7434
7.3267 | -86.7437
7.3275 | -86.743
7.3282 | -86.7435
7.329 | -86.7441
7.3298 | -86.7448
7.3306 | -86.7454
7.3313 | -86.7457
7.3321 | -86.7478
7.3329 | -86.7491
7.3337 | -86.7497
7.3344 | -86.751
7.3352 | -86.7513
7.336 | -86.7522
7.3368 | -86.7532
7.3375 | -86.7546
7.3383 | -86.7556
7.3391 | -86.757
7.3399 | -86.7575
7.3406 | -86.7583
7.3414 | -86.7575
7.3422 | -86.7558
7.343 | -86.7559
7.3437 | -86.7553
7.3445 | -86.752
7.3453 | -86.7529
7.3461 | -86.7536
7.3468 | -86.7534
7.3476 | -86.7536
7.3484 | -86.7554
7.3492 | -86.7562
7.3499 | -86.7571
7.3507 | -86.7593
7.3515 | -86.7618
7.3523 | -86.7631
7.353 | -86.7661
7.3538 | -86.7683
7.3546 | -86.7694
7.3554 | -86.7704
7.3561 | -86.7726
7.3569 | -86.7738
7.3577 | -86.775
7.3585 | -86.7789
7.3592 | -86.7812
7.36 | -86.7814
7.3608 | -86.7809
7.3616 | -86.7804
7.3623 | -86.7821
7.3631 | -86.7831
7.3639 | -86.7835
7.3647 | -86.7842
7.3654 | -86.7855
7.3662 | -86.7877
7.367 | -86.7886
7.3678 | -86.7878
7.3685 | -86.7872
7.3693 | -86.7879
7.3701 | -86.7885
7.3709 | -86.7897
7.3716 | -86.7905
7.3724 | -86.7908
7.3732 | -86.7908
7.3739 | -86.7915
7.3747 | -86.7925
7.3755 | -86.7935
7.3763 | -86.7956
7.377 | -86.7973
7.3778 | -86.8
7.3786 | -86.8014
7.3794 | -86.8012
7.3801 | -86.8017
7.3809 | -86.8028
7.3817 | -86.8037
7.3825 | -86.8055
7.3832 | -86.8087
7.384 | -86.8108
7.3848 | -86.8117
7.3856 | -86.8115
7.3863 | -86.8105
7.3871 | -86.8123
7.3879 | -86.8151
7.3887 | -86.8164
7.3894 | -86.8178
7.3902 | -86.8193
7.391 | -86.8209
7.3918 | -86.8226
7.3925 | -86.8244
7.3933 | -86.8231
7.3941 | -86.8201
7.3949 | -86.819
7.3956 | -86.817
7.3964 | -86.8175
7.3972 | -86.8169
7.398 | -86.8169
7.3987 | -86.8183
7.3995 | -86.8193
7.4003 | -86.8198
7.4011 | -86.8203
7.4018 | -86.8208
7.4026 | -86.8208
7.4034 | -86.8213
7.4042 | -86.8216
7.4049 | -86.8218
7.4057 | -86.8226
7.4065 | -86.825
7.4073 | -86.8261
7.408 | -86.8269
7.4088 | -86.8278
7.4096 | -86.8289
7.4104 | -86.8297
7.4111 | -86.8314
7.4119 | -86.8323
7.4127 | -86.8322
7.4135 | -86.8333
7.4142 | -86.8344
7.415 | -86.8349
7.4158 | -86.8356
7.4166 | -86.836
7.4173 | -86.8365
7.4181 | -86.8381
7.4189 | -86.8396
7.4197 | -86.8417
7.4204 | -86.8437
7.4212 | -86.8446
7.422 | -86.8468
7.4228 | -86.8476
7.4235 | -86.8482
7.4243 | -86.8496
7.4251 | -86.8506
7.4259 | -86.8518
7.4266 | -86.8519
7.4274 | -86.8517
7.4282 | -86.8525
7.429 | -86.854
7.4297 | -86.8549
7.4305 | -86.8561
7.4313 | -86.8574
7.4321 | -86.8579
7.4328 | -86.8601
7.4336 | -86.862
7.4344 | -86.8627
7.4352 | -86.8633
7.4359 | -86.8607
7.4367 | -86.8608
7.4375 | -86.8617
7.4383 | -86.8639
7.439 | -86.8662
7.4398 | -86.8672
7.4406 | -86.8692
7.4414 | -86.8714
7.4421 | -86.8726
7.4429 | -86.8739
7.4437 | -86.8746
7.4445 | -86.8757
7.4452 | -86.8759
7.446 | -86.8745
7.4468 | -86.8751
7.4475 | -86.8764
7.4483 | -86.8773
7.4491 | -86.8776
7.4499 | -86.8785
7.4506 | -86.8797
7.4514 | -86.8808
7.4522 | -86.8842
7.453 | -86.8847
7.4537 | -86.8827
7.4545 | -86.8808
7.4553 | -86.881
7.4561 | -86.8805
7.4568 | -86.8801
7.4576 | -86.8807
7.4584 | -86.8813
7.4592 | -86.8825
7.4599 | -86.8835
7.4607 | -86.8834
7.4615 | -86.8822
7.4623 | -86.8819
7.463 | -86.8819
7.4638 | -86.8828
7.4646 | -86.8836
7.4654 | -86.8851
7.4661 | -86.8859
7.4669 | -86.8871
7.4677 | -86.8885
7.4685 | -86.8907
7.4692 | -86.8923
7.47 | -86.8942
7.4708 | -86.8967
7.4716 | -86.8969
7.4723 | -86.8971
7.4731 | -86.8971
7.4739 | -86.8975
7.4747 | -86.8962
7.4754 | -86.8972
7.4762 | -86.8981
7.477 | -86.8987
7.4778 | -86.899
7.4785 | -86.9025
7.4793 | -86.9024
7.4801 | -86.8993
7.4809 | -86.8971
7.4816 | -86.897
7.4824 | -86.9012
7.4832 | -86.9065
7.484 | -86.9091
7.4847 | -86.9076
7.4855 | -86.9058
7.4863 | -86.904
7.4871 | -86.9009
7.4878 | -86.8995
7.4886 | -86.9003
7.4894 | -86.9002
7.4902 | -86.8998
7.4909 | -86.8995
7.4917 | -86.899
7.4925 | -86.8989
7.4933 | -86.9004
7.494 | -86.9014
7.4948 | -86.9007
7.4956 | -86.8993
7.4964 | -86.8979
7.4971 | -86.898
7.4979 | -86.8994
7.4987 | -86.9012
7.4995 | -86.9021
7.5002 | -86.9009
7.501 | -86.8996
7.5018 | -86.898
7.5026 | -86.8959
7.5033 | -86.894
7.5041 | -86.8923
7.5049 | -86.8921
7.5057 | -86.8922
7.5064 | -86.8925
7.5072 | -86.8906
7.508 | -86.8886
7.5088 | -86.886
7.5095 | -86.8842
7.5103 | -86.8818
7.5111 | -86.8815
7.5119 | -86.881
7.5126 | -86.8794
7.5134 | -86.8784
7.5142 | -86.8779
7.515 | -86.8767
7.5157 | -86.8752
7.5165 | -86.8725
7.5173 | -86.8716
7.5181 | -86.8712
7.5188 | -86.8702
7.5196 | -86.869
7.5204 | -86.8682
7.5211 | -86.8658
7.5219 | -86.8658
7.5227 | -86.8644
7.5235 | -86.864
7.5242 | -86.8653
7.525 | -86.8682
7.5258 | -86.8676
7.5266 | -86.867
7.5273 | -86.8656
7.5281 | -86.8646
7.5289 | -86.8649
7.5297 | -86.8655
7.5304 | -86.8653
7.5312 | -86.8644
7.532 | -86.8632
7.5328 | -86.8614
7.5335 | -86.8598
7.5343 | -86.8589
7.5351 | -86.8578
7.5359 | -86.8574
7.5366 | -86.8566
7.5374 | -86.8568
7.5382 | -86.8558
7.539 | -86.8537
7.5397 | -86.8535
7.5405 | -86.8526
7.5413 | -86.8506
7.5421 | -86.8472
7.5428 | -86.8439
7.5436 | -86.8433
7.5444 | -86.8401
7.5452 | -86.839
7.5459 | -86.8386
7.5467 | -86.8383
7.5475 | -86.8374
7.5483 | -86.8368
7.549 | -86.8351
7.5498 | -86.8333
7.5506 | -86.8306
7.5514 | -86.8281
7.5521 | -86.827
7.5529 | -86.8254
7.5537 | -86.824
7.5545 | -86.8237
7.5552 | -86.8224
7.556 | -86.8199
7.5568 | -86.8167
7.5576 | -86.8151
7.5583 | -86.8141
7.5591 | -86.8129
7.5599 | -86.8122
7.5607 | -86.8116
7.5614 | -86.8099
7.5622 | -86.8081
7.563 | -86.8073
7.5638 | -86.8055
7.5645 | -86.803
7.5653 | -86.8015
7.5661 | -86.7998
7.5669 | -86.798
7.5676 | -86.7969
7.5684 | -86.7949
7.5692 | -86.7927
7.57 | -86.7896
7.5707 | -86.7854
7.5715 | -86.7825
7.5723 | -86.7801
7.5731 | -86.7776
7.5738 | -86.7765
7.5746 | -86.7746
7.5754 | -86.7714
7.5762 | -86.7686
7.5769 | -86.7669
7.5777 | -86.7646
7.5785 | -86.7623
7.5793 | -86.7606
7.58 | -86.7583
7.5808 | -86.7578
7.5816 | -86.7597
7.5824 | -86.7586
7.5831 | -86.7569
7.5839 | -86.7555
7.5847 | -86.7537
7.5855 | -86.7511
7.5862 | -86.748
7.587 | -86.746
7.5878 | -86.7435
7.5886 | -86.7417
7.5893 | -86.7406
7.5901 | -86.7395
7.5909 | -86.7374
7.5917 | -86.7361
7.5924 | -86.7343
7.5932 | -86.7319
7.594 | -86.7294
7.5948 | -86.7276
7.5955 | -86.726
7.5963 | -86.7242
7.5971 | -86.7227
7.5978 | -86.7211
7.5986 | -86.7194
7.5994 | -86.717
7.6002 | -86.7133
7.6009 | -86.7111
7.6017 | -86.7096
7.6025 | -86.7074
7.6033 | -86.7036
7.604 | -86.6995
7.6048 | -86.6975
7.6056 | -86.6955
7.6064 | -86.693
7.6071 | -86.6909
7.6079 | -86.6885
7.6087 | -86.6869
7.6095 | -86.6848
7.6102 | -86.6824
7.611 | -86.6811
7.6118 | -86.6794
7.6126 | -86.6778
7.6133 | -86.6766
7.6141 | -86.6758
7.6149 | -86.6741
7.6157 | -86.6706
7.6164 | -86.6666
7.6172 | -86.6652
7.618 | -86.6631
7.6188 | -86.6606
7.6195 | -86.6587
7.6203 | -86.6564
7.6211 | -86.654
7.6219 | -86.6509
7.6226 | -86.6467
7.6234 | -86.643
7.6242 | -86.6401
7.625 | -86.6377
7.6257 | -86.6354
7.6265 | -86.6332
7.6273 | -86.6313
7.6281 | -86.6294
7.6288 | -86.6269
7.6296 | -86.625
7.6304 | -86.6233
7.6312 | -86.6207
7.6319 | -86.6187
7.6327 | -86.6177
7.6335 | -86.6158
7.6343 | -86.6134
7.635 | -86.6113
7.6358 | -86.6107
7.6366 | -86.6093
7.6374 | -86.6073
7.6381 | -86.6049
7.6389 | -86.6018
7.6397 | -86.6006
7.6405 | -86.5983
7.6412 | -86.5953
7.642 | -86.5934
7.6428 | -86.5911
7.6436 | -86.5885
7.6443 | -86.5873
7.6451 | -86.5854
7.6459 | -86.5831
7.6467 | -86.5842
7.6474 | -86.5845
7.6482 | -86.5824
7.649 | -86.5808
7.6498 | -86.58
7.6505 | -86.5786
7.6513 | -86.5765
7.6521 | -86.5746
7.6529 | -86.5741
7.6536 | -86.5732
7.6544 | -86.5718
7.6552 | -86.57
7.656 | -86.5683
7.6567 | -86.5664
7.6575 | -86.5654
7.6583 | -86.5658
7.6591 | -86.5652
7.6598 | -86.5632
7.6606 | -86.5614
7.6614 | -86.5588
7.6622 | -86.556
7.6629 | -86.5555
7.6637 | -86.5545
7.6645 | -86.5521
7.6653 | -86.551
7.666 | -86.5499
7.6668 | -86.5486
7.6676 | -86.5465
7.6684 | -86.5444
7.6691 | -86.5425
7.6699 | -86.541
7.6707 | -86.5398
7.6714 | -86.5389
7.6722 | -86.5379
7.673 | -86.5365
7.6738 | -86.5354
7.6745 | -86.5345
7.6753 | -86.5342
7.6761 | -86.5338
7.6769 | -86.5321
7.6776 | -86.5298
7.6784 | -86.529
7.6792 | -86.5282
7.68 | -86.5269
7.6807 | -86.5252
7.6815 | -86.5238
7.6823 | -86.5226
7.6831 | -86.5214
7.6838 | -86.5201
7.6846 | -86.519
7.6854 | -86.5181
7.6862 | -86.5174
7.6869 | -86.5163
7.6877 | -86.5147
7.6885 | -86.5136
7.6893 | -86.5152
7.69 | -86.5142
7.6908 | -86.5131
7.6916 | -86.5113
7.6924 | -86.5094
7.6931 | -86.5083
7.6939 | -86.5082
7.6947 | -86.5071
7.6955 | -86.5062
7.6962 | -86.505
7.697 | -86.5043
7.6978 | -86.5035
7.6986 | -86.5013
7.6993 | -86.4994
7.7001 | -86.4988
7.7009 | -86.4979
7.7017 | -86.4968
7.7024 | -86.4963
7.7032 | -86.4958
7.704 | -86.4949
7.7048 | -86.495
7.7055 | -86.4946
7.7063 | -86.4944
7.7071 | -86.4952
7.7079 | -86.4966
7.7086 | -86.4966
7.7094 | -86.4968
7.7102 | -86.4969
7.711 | -86.4966
7.7117 | -86.4986
7.7125 | -86.498
7.7133 | -86.4981
7.7141 | -86.4987
7.7148 | -86.4978
7.7156 | -86.4965
7.7164 | -86.4995
7.7172 | -86.5037
7.7179 | -86.5061
7.7187 | -86.5077
7.7195 | -86.5081
7.7203 | -86.5079
7.721 | -86.5077
7.7218 | -86.5071
7.7226 | -86.5077
7.7234 | -86.5079
7.7241 | -86.5079
7.7249 | -86.5081
7.7257 | -86.508
7.7265 | -86.5083
7.7272 | -86.5087
7.728 | -86.5089
7.7288 | -86.5097
7.7296 | -86.5098
7.7303 | -86.5097
7.7311 | -86.5103
7.7319 | -86.5098
7.7327 | -86.509
7.7334 | -86.5097
7.7342 | -86.5095
7.735 | -86.5093
7.7358 | -86.5099
7.7365 | -86.5125
7.7373 | -86.5139
7.7381 | -86.5146
7.7389 | -86.5151
7.7396 | -86.5146
7.7404 | -86.5149
7.7412 | -86.5157
7.742 | -86.5163
7.7427 | -86.5164
7.7435 | -86.5166
7.7443 | -86.5171
7.745 | -86.5176
7.7458 | -86.5189
7.7466 | -86.5216
7.7474 | -86.5236
7.7481 | -86.5248
7.7489 | -86.5266
7.7497 | -86.528
7.7505 | -86.5296
7.7512 | -86.5314
7.752 | -86.5334
7.7528 | -86.5349
7.7536 | -86.5364
7.7543 | -86.5378
7.7551 | -86.5394
7.7559 | -86.5418
7.7567 | -86.5438
7.7574 | -86.5464
7.7582 | -86.5482
7.759 | -86.5503
7.7598 | -86.5535
7.7605 | -86.5571
7.7613 | -86.5604
7.7621 | -86.5625
7.7629 | -86.5641
7.7636 | -86.5656
7.7644 | -86.5664
7.7652 | -86.5677
7.766 | -86.5706
7.7667 | -86.5741
7.7675 | -86.5757
7.7683 | -86.5764
7.7691 | -86.5784
7.7698 | -86.5809
7.7706 | -86.5828
7.7714 | -86.5839
7.7722 | -86.5863
7.7729 | -86.5892
7.7737 | -86.5924
7.7745 | -86.5945
7.7753 | -86.5978
7.776 | -86.5999
7.7768 | -86.5995
7.7776 | -86.6
7.7784 | -86.6046
7.7791 | -86.6072
7.7799 | -86.6099
7.7807 | -86.6124
7.7815 | -86.615
7.7822 | -86.619
7.783 | -86.6234
7.7838 | -86.628
7.7846 | -86.6311
7.7853 | -86.6339
7.7861 | -86.6362
7.7869 | -86.6386
7.7877 | -86.64
7.7884 | -86.6426
7.7892 | -86.6465
7.79 | -86.6501
7.7908 | -86.6537
7.7915 | -86.6573
7.7923 | -86.6609
7.7931 | -86.6634
7.7939 | -86.6653
7.7946 | -86.6678
7.7954 | -86.6715
7.7962 | -86.6753
7.797 | -86.6787
7.7977 | -86.6816
7.7985 | -86.6849
7.7993 | -86.688
7.8001 | -86.69
7.8008 | -86.6938
7.8016 | -86.698
7.8024 | -86.7041
7.8032 | -86.7072
7.8039 | -86.7099
7.8047 | -86.7142
7.8055 | -86.7189
7.8063 | -86.7222
7.807 | -86.7251
7.8078 | -86.7283
7.8086 | -86.7312
7.8094 | -86.7346
7.8101 | -86.7381
7.8109 | -86.7413
7.8117 | -86.7455
7.8125 | -86.7493
7.8132 | -86.7527
7.814 | -86.7562
7.8148 | -86.7589
7.8156 | -86.7616
7.8163 | -86.765
7.8171 | -86.7687
7.8179 | -86.7725
7.8186 | -86.7764
7.8194 | -86.78
7.8202 | -86.785
7.821 | -86.7899
7.8217 | -86.7934
7.8225 | -86.7961
7.8233 | -86.7999
7.8241 | -86.8035
7.8248 | -86.8073
7.8256 | -86.8109
7.8264 | -86.8144
7.8272 | -86.8189
7.8279 | -86.8226
7.8287 | -86.8261
7.8295 | -86.8302
7.8303 | -86.8345
7.831 | -86.8385
7.8318 | -86.8425
7.8326 | -86.8462
7.8334 | -86.8492
7.8341 | -86.8522
7.8349 | -86.8551
7.8357 | -86.8576
7.8365 | -86.8615
7.8372 | -86.8661
7.838 | -86.8704
7.8388 | -86.8746
7.8396 | -86.8789
7.8403 | -86.8837
7.8411 | -86.8886
7.8419 | -86.893
7.8427 | -86.899
7.8434 | -86.9063
7.8442 | -86.912
7.845 | -86.9158
7.8458 | -86.9196
7.8465 | -86.9241
7.8473 | -86.9284
7.8481 | -86.9324
7.8489 | -86.9368
7.8496 | -86.9408
7.8504 | -86.9461
7.8512 | -86.951
7.852 | -86.9555
7.8527 | -86.96
7.8535 | -86.9641
7.8543 | -86.9689
7.8551 | -86.9748
7.8558 | -86.9791
7.8566 | -86.9839
7.8574 | -86.9887
7.8582 | -86.9935
7.8589 | -86.9982
7.8597 | -87.004
7.8605 | -87.0076
7.8613 | -87.0115
7.862 | -87.0151
7.8628 | -87.0189
7.8636 | -87.0242
7.8644 | -87.0289
7.8651 | -87.0337
7.8659 | -87.0384
7.8667 | -87.0431
7.8675 | -87.0477
7.8682 | -87.0531
7.869 | -87.0574
7.8698 | -87.0619
7.8706 | -87.0676
7.8713 | -87.0745
7.8721 | -87.0807
7.8729 | -87.0873
7.8737 | -87.0958
7.8744 | -87.1014
7.8752 | -87.1057
7.876 | -87.1111
7.8768 | -87.1164
7.8775 | -87.1199
7.8783 | -87.1228
7.8791 | -87.1261
7.8799 | -87.1309
7.8806 | -87.1353
7.8814 | -87.1402
7.8822 | -87.1437
7.883 | -87.1479
7.8837 | -87.1521
7.8845 | -87.1566
7.8853 | -87.161
7.8861 | -87.1654
7.8868 | -87.1716
7.8876 | -87.1775
7.8884 | -87.1821
7.8892 | -87.1868
7.8899 | -87.1918
7.8907 | -87.1967
7.8915 | -87.2028
7.8922 | -87.208
7.893 | -87.2124
7.8938 | -87.2168
7.8946 | -87.2218
7.8953 | -87.2265
7.8961 | -87.2314
7.8969 | -87.2355
7.8977 | -87.2404
7.8984 | -87.2466
7.8992 | -87.2529
7.9 | -87.2575
7.9008 | -87.2624
7.9015 | -87.2671
7.9023 | -87.2719
7.9031 | -87.2775
7.9039 | -87.284
7.9046 | -87.2889
7.9054 | -87.2933
7.9062 | -87.2979
7.907 | -87.3027
7.9077 | -87.3077
7.9085 | -87.3119
7.9093 | -87.3161
7.9101 | -87.321
7.9108 | -87.3266
7.9116 | -87.3317
7.9124 | -87.3348
7.9132 | -87.3386
7.9139 | -87.3434
7.9147 | -87.3477
7.9155 | -87.3519
7.9163 | -87.3563
7.917 | -87.3599
7.9178 | -87.365
7.9186 | -87.3709
7.9194 | -87.3759
7.9201 | -87.3801
7.9209 | -87.3834
7.9217 | -87.3882
7.9225 | -87.3927
7.9232 | -87.3962
7.924 | -87.4007
7.9248 | -87.4063
7.9256 | -87.4135
7.9263 | -87.4221
7.9271 | -87.4286
7.9279 | -87.4348
7.9287 | -87.4382
7.9294 | -87.4438
7.9302 | -87.449
7.931 | -87.4534
7.9318 | -87.458
7.9325 | -87.4623
7.9333 | -87.4669
7.9341 | -87.4718
7.9349 | -87.4765
7.9356 | -87.4825
7.9364 | -87.4887
7.9372 | -87.4932
7.938 | -87.4972
7.9387 | -87.5019
7.9395 | -87.5063
7.9403 | -87.5102
7.9411 | -87.5144
7.9418 | -87.5188
7.9426 | -87.523
7.9434 | -87.5273
7.9442 | -87.5322
7.9449 | -87.5374
7.9457 | -87.5419
7.9465 | -87.548
7.9473 | -87.5538
7.948 | -87.558
7.9488 | -87.5623
7.9496 | -87.567
7.9504 | -87.572
7.9511 | -87.5766
7.9519 | -87.5809
7.9527 | -87.5847
7.9535 | -87.5895
7.9542 | -87.5949
7.955 | -87.5995
7.9558 | -87.6045
7.9566 | -87.6072
7.9573 | -87.6101
7.9581 | -87.6149
7.9589 | -87.6195
7.9597 | -87.6238
7.9604 | -87.6275
7.9612 | -87.6312
7.962 | -87.6355
7.9628 | -87.639
7.9635 | -87.6426
7.9643 | -87.6467
7.9651 | -87.6512
7.9658 | -87.6558
7.9666 | -87.6598
7.9674 | -87.6644
7.9682 | -87.6694
7.9689 | -87.6739
7.9697 | -87.6794
7.9705 | -87.685
7.9713 | -87.6931
7.972 | -87.698
7.9728 | -87.701
7.9736 | -87.7041
7.9744 | -87.7074
7.9751 | -87.7105
7.9759 | -87.7139
7.9767 | -87.7181
7.9775 | -87.7229
7.9782 | -87.7276
7.979 | -87.7321
7.9798 | -87.7354
7.9806 | -87.7401
7.9813 | -87.7451
7.9821 | -87.7514
7.9829 | -87.7585
7.9837 | -87.763
7.9844 | -87.7659
7.9852 | -87.7689
7.986 | -87.7721
7.9868 | -87.7743
7.9875 | -87.7783
7.9883 | -87.7829
7.9891 | -87.787
7.9899 | -87.7921
7.9906 | -87.7988
7.9914 | -87.806
7.9922 | -87.8113
7.993 | -87.8163
7.9937 | -87.8208
7.9945 | -87.8255
7.9953 | -87.8311
7.9961 | -87.8348
7.9968 | -87.8386
7.9976 | -87.8433
7.9984 | -87.8481
7.9992 | -87.8525
7.9999 | -87.8559
8.0007 | -87.8591
8.0015 | -87.8625
8.0023 | -87.8674
8.003 | -87.8709
8.0038 | -87.8738
8.0046 | -87.8779
8.0054 | -87.8826
8.0061 | -87.8868
8.0069 | -87.8904
8.0077 | -87.8945
8.0085 | -87.8992
8.0092 | -87.9052
8.01 | -87.9095
8.0108 | -87.9141
8.0116 | -87.919
8.0123 | -87.9235
8.0131 | -87.9295
8.0139 | -87.9352
8.0147 | -87.939
8.0154 | -87.9432
8.0162 | -87.9466
8.017 | -87.9509
8.0178 | -87.9582
8.0185 | -87.9661
8.0193 | -87.9735
8.0201 | -87.9784
8.0209 | -87.982
8.0216 | -87.986
8.0224 | -87.9903
8.0232 | -87.996
8.024 | -88.002
8.0247 | -88.0066
8.0255 | -88.0133
8.0263 | -88.0219
8.0271 | -88.0277
8.0278 | -88.0301
8.0286 | -88.0302
8.0294 | -88.0313
8.0302 | -88.034
8.0309 | -88.0377
8.0317 | -88.0419
8.0325 | -88.0457
8.0333 | -88.0492
8.034 | -88.0534
8.0348 | -88.0592
8.0356 | -88.0648
8.0364 | -88.0701
8.0371 | -88.0748
8.0379 | -88.0788
8.0387 | -88.0833
8.0395 | -88.088
8.0402 | -88.0929
8.041 | -88.0972
8.0418 | -88.1025
8.0425 | -88.1087
8.0433 | -88.1141
8.0441 | -88.1196
8.0449 | -88.1226
8.0456 | -88.1239
8.0464 | -88.1232
8.0472 | -88.1262
8.048 | -88.1305
8.0487 | -88.1354
8.0495 | -88.1411
8.0503 | -88.1461
8.0511 | -88.1502
8.0518 | -88.1555
8.0526 | -88.1592
8.0534 | -88.1634
8.0542 | -88.1684
8.0549 | -88.1715
8.0557 | -88.1732
8.0565 | -88.1764
8.0573 | -88.1813
8.058 | -88.1884
8.0588 | -88.1921
8.0596 | -88.1956
8.0604 | -88.1994
8.0611 | -88.2031
8.0619 | -88.2068
8.0627 | -88.2127
8.0635 | -88.2161
8.0642 | -88.219
8.065 | -88.2206
8.0658 | -88.2244
8.0666 | -88.2275
8.0673 | -88.2313
8.0681 | -88.2341
8.0689 | -88.2375
8.0697 | -88.2419
8.0704 | -88.246
8.0712 | -88.2493
8.072 | -88.2531
8.0728 | -88.2565
8.0735 | -88.2595
8.0743 | -88.2633
8.0751 | -88.2667
8.0759 | -88.2721
8.0766 | -88.2766
8.0774 | -88.2803
8.0782 | -88.2841
8.079 | -88.2878
8.0797 | -88.291
8.0805 | -88.2947
8.0813 | -88.2984
8.0821 | -88.3033
8.0828 | -88.3071
8.0836 | -88.3116
8.0844 | -88.317
8.0852 | -88.3246
8.0859 | -88.3316
8.0867 | -88.3351
8.0875 | -88.3417
8.0883 | -88.347
8.089 | -88.3506
8.0898 | -88.3534
8.0906 | -88.3562
8.0914 | -88.3605
8.0921 | -88.3629
8.0929 | -88.3658
8.0937 | -88.3692
8.0945 | -88.3743
8.0952 | -88.3788
8.096 | -88.381
8.0968 | -88.3847
8.0976 | -88.3885
8.0983 | -88.3921
8.0991 | -88.3955
8.0999 | -88.3976
8.1007 | -88.4003
8.1014 | -88.4042
8.1022 | -88.4067
8.103 | -88.4086
8.1038 | -88.4115
8.1045 | -88.4146
8.1053 | -88.4177
8.1061 | -88.4215
8.1069 | -88.4249
8.1076 | -88.4286
8.1084 | -88.433
8.1092 | -88.438
8.11 | -88.4419
8.1107 | -88.4461
8.1115 | -88.4504
8.1123 | -88.4536
8.113 | -88.4572
8.1138 | -88.4593
8.1146 | -88.462
8.1154 | -88.4659
8.1161 | -88.4719
8.1169 | -88.4767
8.1177 | -88.4799
8.1185 | -88.4838
8.1192 | -88.4868
8.12 | -88.4903
8.1208 | -88.4944
8.1216 | -88.4983
8.1223 | -88.5015
8.1231 | -88.5039
8.1239 | -88.507
8.1247 | -88.5104
8.1254 | -88.5136
8.1262 | -88.5168
8.127 | -88.519
8.1278 | -88.5213
8.1285 | -88.5247
8.1293 | -88.5286
8.1301 | -88.5315
8.1309 | -88.535
8.1316 | -88.5393
8.1324 | -88.5438
8.1332 | -88.5472
8.134 | -88.5511
8.1347 | -88.5554
8.1355 | -88.5595
8.1363 | -88.5624
8.1371 | -88.565
8.1378 | -88.5684
8.1386 | -88.5717
8.1394 | -88.575
8.1402 | -88.5783
8.1409 | -88.5817
8.1417 | -88.5857
8.1425 | -88.5904
8.1433 | -88.5939
8.144 | -88.5967
8.1448 | -88.5975
8.1456 | -88.5996
8.1464 | -88.6029
8.1471 | -88.6056
8.1479 | -88.6107
8.1487 | -88.6141
8.1495 | -88.617
8.1502 | -88.6194
8.151 | -88.6247
8.1518 | -88.63
8.1526 | -88.633
8.1533 | -88.6354
8.1541 | -88.639
8.1549 | -88.6416
8.1557 | -88.6447
8.1564 | -88.6474
8.1572 | -88.6501
8.158 | -88.653
8.1588 | -88.655
8.1595 | -88.6574
8.1603 | -88.6604
8.1611 | -88.6638
8.1619 | -88.6678
8.1626 | -88.6705
8.1634 | -88.6728
8.1642 | -88.6777
8.165 | -88.6844
8.1657 | -88.6883
8.1665 | -88.6917
8.1673 | -88.6944
8.1681 | -88.6966
8.1688 | -88.698
8.1696 | -88.7003
8.1704 | -88.7034
8.1712 | -88.7075
8.1719 | -88.7135
8.1727 | -88.7196
8.1735 | -88.7235
8.1743 | -88.7259
8.175 | -88.7282
8.1758 | -88.7295
8.1766 | -88.7315
8.1774 | -88.7336
8.1781 | -88.736
8.1789 | -88.7382
8.1797 | -88.7405
8.1805 | -88.7432
8.1812 | -88.7456
8.182 | -88.749
8.1828 | -88.7525
8.1836 | -88.7539
8.1843 | -88.7559
8.1851 | -88.758
8.1859 | -88.7594
8.1867 | -88.7612
8.1874 | -88.7637
8.1882 | -88.7669
8.189 | -88.7698
8.1897 | -88.7724
8.1905 | -88.775
8.1913 | -88.7767
8.1921 | -88.7773
8.1928 | -88.7781
8.1936 | -88.7807
8.1944 | -88.7817
8.1952 | -88.7838
8.1959 | -88.786
8.1967 | -88.7877
8.1975 | -88.7898
8.1983 | -88.7912
8.199 | -88.7933
8.1998 | -88.7951
8.2006 | -88.7972
8.2014 | -88.8005
8.2021 | -88.8028
8.2029 | -88.8035
8.2037 | -88.8053
8.2045 | -88.8076
8.2052 | -88.8098
8.206 | -88.811
8.2068 | -88.8124
8.2076 | -88.8148
8.2083 | -88.8168
8.2091 | -88.8197
8.2099 | -88.8217
8.2107 | -88.822
8.2114 | -88.8234
8.2122 | -88.8256
8.213 | -88.8302
8.2138 | -88.8334
8.2145 | -88.8325
8.2153 | -88.833
8.2161 | -88.8354
8.2169 | -88.8377
8.2176 | -88.8391
8.2184 | -88.8404
8.2192 | -88.8441
8.22 | -88.8467
8.2207 | -88.8495
8.2215 | -88.8515
8.2223 | -88.8538
8.2231 | -88.8564
8.2238 | -88.8588
8.2246 | -88.8604
8.2254 | -88.8609
8.2262 | -88.862
8.2269 | -88.8627
8.2277 | -88.8627
8.2285 | -88.8631
8.2293 | -88.8643
8.23 | -88.8672
8.2308 | -88.8707
8.2316 | -88.8729
8.2324 | -88.8745
8.2331 | -88.8769
8.2339 | -88.879
8.2347 | -88.8813
8.2355 | -88.8796
8.2362 | -88.8795
8.237 | -88.8808
8.2378 | -88.8822
8.2386 | -88.8841
8.2393 | -88.8864
8.2401 | -88.887
8.2409 | -88.8885
8.2417 | -88.8897
8.2424 | -88.8912
8.2432 | -88.8921
8.244 | -88.8919
8.2448 | -88.8932
8.2455 | -88.8966
8.2463 | -88.9002
8.2471 | -88.9028
8.2479 | -88.905
8.2486 | -88.9056
8.2494 | -88.9077
8.2502 | -88.9095
8.251 | -88.911
8.2517 | -88.9122
8.2525 | -88.9151
8.2533 | -88.9192
8.2541 | -88.9244
8.2548 | -88.9283
8.2556 | -88.9311
8.2564 | -88.9334
8.2572 | -88.9347
8.2579 | -88.9362
8.2587 | -88.9378
8.2595 | -88.9401
8.2602 | -88.9421
8.261 | -88.9432
8.2618 | -88.9438
8.2626 | -88.9446
8.2633 | -88.9461
8.2641 | -88.9472
8.2649 | -88.9489
8.2657 | -88.9507
8.2664 | -88.9527
8.2672 | -88.9549
8.268 | -88.9566
8.2688 | -88.9589
8.2695 | -88.9599
8.2703 | -88.9606
8.2711 | -88.9639
8.2719 | -88.9684
8.2726 | -88.9699
8.2734 | -88.9709
8.2742 | -88.9707
8.275 | -88.9724
8.2757 | -88.9729
8.2765 | -88.9737
8.2773 | -88.9752
8.2781 | -88.9787
8.2788 | -88.9806
8.2796 | -88.982
8.2804 | -88.9839
8.2812 | -88.9857
8.2819 | -88.9884
8.2827 | -88.9897
8.2835 | -88.9902
8.2843 | -88.9902
8.285 | -88.9912
8.2858 | -88.9922
8.2866 | -88.9941
8.2874 | -88.9956
8.2881 | -88.9986
8.2889 | -89.0033
8.2897 | -89.0074
8.2905 | -89.0105
8.2912 | -89.0141
8.292 | -89.0167
8.2928 | -89.0174
8.2936 | -89.02
8.2943 | -89.0206
8.2951 | -89.0209
8.2959 | -89.0213
8.2967 | -89.0239
8.2974 | -89.0275
8.2982 | -89.0297
8.299 | -89.0306
8.2998 | -89.0321
8.3005 | -89.0353
8.3013 | -89.0383
8.3021 | -89.0404
8.3029 | -89.0425
8.3036 | -89.0422
8.3044 | -89.0414
8.3052 | -89.0421
8.306 | -89.0436
8.3067 | -89.0441
8.3075 | -89.0448
8.3083 | -89.0466
8.3091 | -89.0478
8.3098 | -89.0479
8.3106 | -89.0484
8.3114 | -89.0485
8.3122 | -89.0478
8.3129 | -89.0473
8.3137 | -89.0487
8.3145 | -89.0526
8.3153 | -89.0549
8.316 | -89.0564
8.3168 | -89.0565
8.3176 | -89.0565
8.3184 | -89.0571
8.3191 | -89.0566
8.3199 | -89.058
8.3207 | -89.0604
8.3215 | -89.062
8.3222 | -89.0615
8.323 | -89.0628
8.3238 | -89.065
8.3246 | -89.0674
8.3253 | -89.0687
8.3261 | -89.0703
8.3269 | -89.0729
8.3277 | -89.0764
8.3284 | -89.0778
8.3292 | -89.0784
8.33 | -89.0801
8.3308 | -89.0825
8.3315 | -89.0828
8.3323 | -89.0848
8.3331 | -89.0852
8.3338 | -89.0824
8.3346 | -89.0816
8.3354 | -89.0831
8.3362 | -89.0886
8.3369 | -89.0893
8.3377 | -89.0917
8.3385 | -89.0914
8.3393 | -89.093
8.34 | -89.0942
8.3408 | -89.0951
8.3416 | -89.0953
8.3424 | -89.0966
8.3431 | -89.0977
8.3439 | -89.0996
8.3447 | -89.1003
8.3455 | -89.1008
8.3462 | -89.1015
8.347 | -89.1039
8.3478 | -89.1032
8.3486 | -89.1055
8.3493 | -89.1063
8.3501 | -89.1018
8.3509 | -89.1017
8.3517 | -89.1015
8.3524 | -89.0994
8.3532 | -89.0996
8.354 | -89.1019
8.3548 | -89.1043
8.3555 | -89.1056
8.3563 | -89.1055
8.3571 | -89.105
8.3579 | -89.1042
8.3586 | -89.1042
8.3594 | -89.1049
8.3602 | -89.1048
8.361 | -89.1021
8.3617 | -89.1011
8.3625 | -89.1029
8.3633 | -89.104
8.3641 | -89.1048
8.3648 | -89.1042
8.3656 | -89.1036
8.3664 | -89.1048
8.3672 | -89.1044
8.3679 | -89.1051
8.3687 | -89.1063
8.3695 | -89.1061
8.3703 | -89.1073
8.371 | -89.1084
8.3718 | -89.1083
8.3726 | -89.1073
8.3734 | -89.107
8.3741 | -89.1066
8.3749 | -89.1064
8.3757 | -89.1056
8.3765 | -89.1048
8.3772 | -89.1064
8.378 | -89.1084
8.3788 | -89.1097
8.3796 | -89.1101
8.3803 | -89.1101
8.3811 | -89.1102
8.3819 | -89.1103
8.3827 | -89.1113
8.3834 | -89.1128
8.3842 | -89.1137
8.385 | -89.1153
8.3858 | -89.1166
8.3865 | -89.1174
8.3873 | -89.119
8.3881 | -89.1191
8.3889 | -89.1183
8.3896 | -89.1192
8.3904 | -89.1196
8.3912 | -89.1196
8.392 | -89.1201
8.3927 | -89.1203
8.3935 | -89.1202
8.3943 | -89.12
8.3951 | -89.1201
8.3958 | -89.1184
8.3966 | -89.1174
8.3974 | -89.1167
8.3982 | -89.1171
8.3989 | -89.1185
8.3997 | -89.1209
8.4005 | -89.1219
8.4013 | -89.1221
8.402 | -89.1217
8.4028 | -89.1218
8.4036 | -89.1211
8.4044 | -89.1191
8.4051 | -89.1188
8.4059 | -89.1187
8.4067 | -89.118
8.4075 | -89.1177
8.4082 | -89.1183
8.409 | -89.1182
8.4098 | -89.1187
8.4105 | -89.1202
8.4113 | -89.1207
8.4121 | -89.1203
8.4129 | -89.1201
8.4136 | -89.12
8.4144 | -89.12
8.4152 | -89.1197
8.416 | -89.1187
8.4167 | -89.1182
8.4175 | -89.1181
8.4183 | -89.1179
8.4191 | -89.1176
8.4198 | -89.1173
8.4206 | -89.1169
8.4214 | -89.1163
8.4222 | -89.1147
8.4229 | -89.1145
8.4237 | -89.1139
8.4245 | -89.1137
8.4253 | -89.1133
8.426 | -89.114
8.4268 | -89.1127
8.4276 | -89.1131
8.4284 | -89.1143
8.4291 | -89.1153
8.4299 | -89.1156
8.4307 | -89.1154
8.4315 | -89.1143
8.4322 | -89.1131
8.433 | -89.1123
8.4338 | -89.1117
8.4346 | -89.1113
8.4353 | -89.1104
8.4361 | -89.1105
8.4369 | -89.1102
8.4377 | -89.1105
8.4384 | -89.1097
8.4392 | -89.1084
8.44 | -89.1081
8.4408 | -89.1074
8.4415 | -89.1072
8.4423 | -89.1068
8.4431 | -89.1049
8.4439 | -89.1022
8.4446 | -89.1021
8.4454 | -89.1041
8.4462 | -89.1037
8.447 | -89.1026
8.4477 | -89.1023
8.4485 | -89.1012
8.4493 | -89.0995
8.4501 | -89.1007
8.4508 | -89.1001
8.4516 | -89.0996
8.4524 | -89.1003
8.4532 | -89.0998
8.4539 | -89.099
8.4547 | -89.0974
8.4555 | -89.0959
8.4563 | -89.0951
8.457 | -89.094
8.4578 | -89.0933
8.4586 | -89.0929
8.4594 | -89.0923
8.4601 | -89.0913
8.4609 | -89.0909
8.4617 | -89.0903
8.4625 | -89.0895
8.4632 | -89.0874
8.464 | -89.0863
8.4648 | -89.0855
8.4656 | -89.0847
8.4663 | -89.0824
8.4671 | -89.079
8.4679 | -89.0779
8.4687 | -89.0771
8.4694 | -89.0765
8.4702 | -89.075
8.471 | -89.0728
8.4718 | -89.0714
8.4725 | -89.0691
8.4733 | -89.0679
8.4741 | -89.0672
8.4749 | -89.0652
8.4756 | -89.0636
8.4764 | -89.0623
8.4772 | -89.061
8.478 | -89.061
8.4787 | -89.0606
8.4795 | -89.059
8.4803 | -89.0584
8.481 | -89.0571
8.4818 | -89.0549
8.4826 | -89.0543
8.4834 | -89.0546
8.4841 | -89.0511
8.4849 | -89.0471
8.4857 | -89.046
8.4865 | -89.0454
8.4872 | -89.0464
8.488 | -89.0472
8.4888 | -89.0471
8.4896 | -89.0453
8.4903 | -89.0441
8.4911 | -89.0444
8.4919 | -89.0433
8.4927 | -89.0431
8.4934 | -89.0423
8.4942 | -89.0425
8.495 | -89.0414
8.4958 | -89.0401
8.4965 | -89.0388
8.4973 | -89.0368
8.4981 | -89.0353
8.4989 | -89.0359
8.4996 | -89.0358
8.5004 | -89.035
8.5012 | -89.0338
8.502 | -89.0328
8.5027 | -89.0312
8.5035 | -89.0279
8.5043 | -89.0264
8.5051 | -89.025
8.5058 | -89.0233
8.5066 | -89.0219
8.5074 | -89.0217
8.5082 | -89.0207
8.5089 | -89.0153
8.5097 | -89.0132
8.5105 | -89.0123
8.5113 | -89.0105
8.512 | -89.0084
8.5128 | -89.0065
8.5136 | -89.004
8.5144 | -89.002
8.5151 | -89.0001
8.5159 | -88.9992
8.5167 | -88.9972
8.5175 | -88.9955
8.5182 | -88.9933
8.519 | -88.991
8.5198 | -88.989
8.5206 | -88.9876
8.5213 | -88.9866
8.5221 | -88.9846
8.5229 | -88.9828
8.5237 | -88.9813
8.5244 | -88.9797
8.5252 | -88.9787
8.526 | -88.9772
8.5268 | -88.9757
8.5275 | -88.9744
8.5283 | -88.9728
8.5291 | -88.9732
8.5299 | -88.974
8.5306 | -88.9729
8.5314 | -88.971
8.5322 | -88.9698
8.533 | -88.9683
8.5337 | -88.968
8.5345 | -88.9675
8.5353 | -88.9672
8.5361 | -88.9672
8.5368 | -88.9666
8.5376 | -88.9654
8.5384 | -88.9634
8.5392 | -88.9603
8.5399 | -88.9574
8.5407 | -88.9545
8.5415 | -88.9541
8.5423 | -88.9524
8.543 | -88.9512
8.5438 | -88.9497
8.5446 | -88.9471
8.5454 | -88.9435
8.5461 | -88.9415
8.5469 | -88.9411
8.5477 | -88.9408
8.5485 | -88.9404
8.5492 | -88.9402
8.55 | -88.9391
8.5508 | -88.9382
8.5516 | -88.9365
8.5523 | -88.9339
8.5531 | -88.9321
8.5539 | -88.9317
8.5547 | -88.9294
8.5554 | -88.927
8.5562 | -88.9256
8.557 | -88.924
8.5577 | -88.9226
8.5585 | -88.9213
8.5593 | -88.9215
8.5601 | -88.922
8.5608 | -88.921
8.5616 | -88.92
8.5624 | -88.92
8.5632 | -88.9185
8.5639 | -88.9174
8.5647 | -88.9164
8.5655 | -88.9152
8.5663 | -88.9145
8.567 | -88.9135
8.5678 | -88.9126
8.5686 | -88.9116
8.5694 | -88.9107
8.5701 | -88.9102
8.5709 | -88.9085
8.5717 | -88.9076
8.5725 | -88.9063
8.5732 | -88.9053
8.574 | -88.9053
8.5748 | -88.9049
8.5756 | -88.9048
8.5763 | -88.9048
8.5771 | -88.904
8.5779 | -88.9027
8.5787 | -88.9015
8.5794 | -88.901
8.5802 | -88.9009
8.581 | -88.8993
8.5818 | -88.8976
8.5825 | -88.8958
8.5833 | -88.8961
8.5841 | -88.8949
8.5849 | -88.8937
8.5856 | -88.8924
8.5864 | -88.8915
8.5872 | -88.8912
8.588 | -88.8902
8.5887 | -88.8889
8.5895 | -88.8881
8.5903 | -88.8873
8.5911 | -88.8862
8.5918 | -88.8847
8.5926 | -88.8839
8.5934 | -88.8837
8.5942 | -88.8831
8.5949 | -88.8821
8.5957 | -88.8814
8.5965 | -88.8814
8.5973 | -88.8816
8.598 | -88.8812
8.5988 | -88.8805
8.5996 | -88.8802
8.6004 | -88.8792
8.6011 | -88.8767
8.6019 | -88.8753
8.6027 | -88.8745
8.6035 | -88.8739
8.6042 | -88.8732
8.605 | -88.8709
8.6058 | -88.871
8.6066 | -88.8714
8.6073 | -88.8744
8.6081 | -88.8772
8.6089 | -88.8781
8.6097 | -88.8781
8.6104 | -88.8776
8.6112 | -88.8773
8.612 | -88.877
8.6128 | -88.8759
8.6135 | -88.8752
8.6143 | -88.8754
8.6151 | -88.8747
8.6159 | -88.8739
8.6166 | -88.8736
8.6174 | -88.8733
8.6182 | -88.8731
8.619 | -88.8738
8.6197 | -88.8741
8.6205 | -88.8731
8.6213 | -88.8727
8.6221 | -88.8723
8.6228 | -88.8706
8.6236 | -88.8713
8.6244 | -88.8728
8.6252 | -88.8738
8.6259 | -88.8766
8.6267 | -88.8777
8.6275 | -88.879
8.6282 | -88.8803
8.629 | -88.8794
8.6298 | -88.8797
8.6306 | -88.8794
8.6313 | -88.8778
8.6321 | -88.8766
8.6329 | -88.8765
8.6337 | -88.8765
8.6344 | -88.876
8.6352 | -88.8767
8.636 | -88.8785
8.6368 | -88.8788
8.6375 | -88.8788
8.6383 | -88.8787
8.6391 | -88.8793
8.6399 | -88.8788
8.6406 | -88.879
8.6414 | -88.8792
8.6422 | -88.8794
8.643 | -88.8796
8.6437 | -88.8798
8.6445 | -88.8813
8.6453 | -88.8837
8.6461 | -88.884
8.6468 | -88.8847
8.6476 | -88.8847
8.6484 | -88.8856
8.6492 | -88.8864
8.6499 | -88.8888
8.6507 | -88.8894
8.6515 | -88.89
8.6523 | -88.8892
8.653 | -88.8891
8.6538 | -88.8901
8.6546 | -88.8904
8.6554 | -88.8908
8.6561 | -88.8916
8.6569 | -88.8919
8.6577 | -88.8919
8.6585 | -88.893
8.6592 | -88.8935
8.66 | -88.8942
8.6608 | -88.8954
8.6616 | -88.8966
8.6623 | -88.8977
8.6631 | -88.8985
8.6639 | -88.8992
8.6647 | -88.902
8.6654 | -88.9044
8.6662 | -88.9055
8.667 | -88.9065
8.6678 | -88.9072
8.6685 | -88.9076
8.6693 | -88.9081
8.6701 | -88.9084
8.6709 | -88.9089
8.6716 | -88.9108
8.6724 | -88.9133
8.6732 | -88.9155
8.674 | -88.9163
8.6747 | -88.9173
8.6755 | -88.9188
8.6763 | -88.9199
8.6771 | -88.9208
8.6778 | -88.9218
8.6786 | -88.9235
8.6794 | -88.926
8.6802 | -88.927
8.6809 | -88.9291
8.6817 | -88.9307
8.6825 | -88.9324
8.6833 | -88.9347
8.684 | -88.9385
8.6848 | -88.9416
8.6856 | -88.9444
8.6864 | -88.9467
8.6871 | -88.9468
8.6879 | -88.9466
8.6887 | -88.9477
8.6895 | -88.9489
8.6902 | -88.9505
8.691 | -88.952
8.6918 | -88.9536
8.6926 | -88.9543
8.6933 | -88.9557
8.6941 | -88.9568
8.6949 | -88.9594
8.6957 | -88.9613
8.6964 | -88.9629
8.6972 | -88.9647
8.698 | -88.9673
8.6988 | -88.9694
8.6995 | -88.972
8.7003 | -88.9746
8.7011 | -88.977
8.7019 | -88.9787
8.7026 | -88.9811
8.7034 | -88.9825
8.7042 | -88.9818
8.7049 | -88.9825
8.7057 | -88.9853
8.7065 | -88.9874
8.7073 | -88.9888
8.708 | -88.9913
8.7088 | -88.993
8.7096 | -88.9948
8.7104 | -88.997
8.7111 | -88.9987
8.7119 | -89.0018
8.7127 | -89.0044
8.7135 | -89.0058
8.7142 | -89.0077
8.715 | -89.0097
8.7158 | -89.0119
8.7166 | -89.0142
8.7173 | -89.0178
8.7181 | -89.0201
8.7189 | -89.0222
8.7197 | -89.0241
8.7204 | -89.0254
8.7212 | -89.0249
8.722 | -89.0267
8.7228 | -89.0289
8.7235 | -89.0308
8.7243 | -89.0325
8.7251 | -89.034
8.7259 | -89.0356
8.7266 | -89.0379
8.7274 | -89.0398
8.7282 | -89.0408
8.729 | -89.0429
8.7297 | -89.0452
8.7305 | -89.0476
8.7313 | -89.0516
8.7321 | -89.0548
8.7328 | -89.0568
8.7336 | -89.0603
8.7344 | -89.0627
8.7352 | -89.066
8.7359 | -89.0683
8.7367 | -89.0709
8.7375 | -89.073
8.7383 | -89.0757
8.739 | -89.0788
8.7398 | -89.0825
8.7406 | -89.0852
8.7414 | -89.0871
8.7421 | -89.0902
8.7429 | -89.0922
8.7437 | -89.0944
8.7445 | -89.098
8.7452 | -89.1023
8.746 | -89.1056
8.7468 | -89.1085
8.7476 | -89.1115
8.7483 | -89.1134
8.7491 | -89.1154
8.7499 | -89.1174
8.7507 | -89.1199
8.7514 | -89.1226
8.7522 | -89.1247
8.753 | -89.1272
8.7538 | -89.1299
8.7545 | -89.1328
8.7553 | -89.1358
8.7561 | -89.1387
8.7569 | -89.1421
8.7576 | -89.1442
8.7584 | -89.1467
8.7592 | -89.1499
8.76 | -89.1526
8.7607 | -89.1551
8.7615 | -89.1581
8.7623 | -89.1625
8.7631 | -89.1657
8.7638 | -89.1686
8.7646 | -89.1728
8.7654 | -89.1762
8.7662 | -89.1795
8.7669 | -89.1828
8.7677 | -89.1861
8.7685 | -89.1894
8.7693 | -89.1922
8.77 | -89.1934
8.7708 | -89.1969
8.7716 | -89.1991
8.7724 | -89.2016
8.7731 | -89.2056
8.7739 | -89.2092
8.7747 | -89.2127
8.7754 | -89.2167
8.7762 | -89.2203
8.777 | -89.2224
8.7778 | -89.2243
8.7785 | -89.2279
8.7793 | -89.2301
8.7801 | -89.2321
8.7809 | -89.2362
8.7816 | -89.2413
8.7824 | -89.2461
8.7832 | -89.2496
8.784 | -89.2532
8.7847 | -89.2564
8.7855 | -89.2603
8.7863 | -89.2639
8.7871 | -89.268
8.7878 | -89.2705
8.7886 | -89.2742
8.7894 | -89.2774
8.7902 | -89.2804
8.7909 | -89.2841
8.7917 | -89.2868
8.7925 | -89.2899
8.7933 | -89.2937
8.794 | -89.2983
8.7948 | -89.3021
8.7956 | -89.3041
8.7964 | -89.3055
8.7971 | -89.3083
8.7979 | -89.312
8.7987 | -89.315
8.7995 | -89.3177
8.8002 | -89.3207
8.801 | -89.3243
8.8018 | -89.3295
8.8026 | -89.3341
8.8033 | -89.3373
8.8041 | -89.3407
8.8049 | -89.3448
8.8057 | -89.3508
8.8064 | -89.3549
8.8072 | -89.3592
8.808 | -89.3623
8.8088 | -89.3649
8.8095 | -89.3688
8.8103 | -89.3727
8.8111 | -89.3763
8.8119 | -89.3789
8.8126 | -89.3816
8.8134 | -89.3855
8.8142 | -89.389
8.815 | -89.3921
8.8157 | -89.3952
8.8165 | -89.3991
8.8173 | -89.4032
8.8181 | -89.4066
8.8188 | -89.4098
8.8196 | -89.4138
8.8204 | -89.418
8.8212 | -89.4227
8.8219 | -89.4267
8.8227 | -89.4299
8.8235 | -89.4334
8.8243 | -89.4368
8.825 | -89.44
8.8258 | -89.4433
8.8266 | -89.447
8.8274 | -89.4503
8.8281 | -89.457
8.8289 | -89.4612
8.8297 | -89.4663
8.8305 | -89.47
8.8312 | -89.4728
8.832 | -89.4755
8.8328 | -89.4792
8.8336 | -89.4849
8.8343 | -89.4891
8.8351 | -89.4926
8.8359 | -89.4958
8.8367 | -89.4997
8.8374 | -89.5049
8.8382 | -89.5082
8.839 | -89.5114
8.8398 | -89.5142
8.8405 | -89.5173
8.8413 | -89.5203
8.8421 | -89.5234
8.8429 | -89.5275
8.8436 | -89.5308
8.8444 | -89.5342
8.8452 | -89.5384
8.846 | -89.5423
8.8467 | -89.5454
8.8475 | -89.5487
8.8483 | -89.5522
8.8491 | -89.5555
8.8498 | -89.5584
8.8506 | -89.5617
8.8514 | -89.5643
8.8521 | -89.5676
8.8529 | -89.5719
8.8537 | -89.5759
8.8545 | -89.5794
8.8552 | -89.5831
8.856 | -89.5869
8.8568 | -89.5907
8.8576 | -89.5951
8.8583 | -89.5981
8.8591 | -89.6015
8.8599 | -89.6036
8.8607 | -89.6069
8.8614 | -89.6099
8.8622 | -89.6133
8.863 | -89.6175
8.8638 | -89.6216
8.8645 | -89.6252
8.8653 | -89.628
8.8661 | -89.6331
8.8669 | -89.6367
8.8676 | -89.6398
8.8684 | -89.6435
8.8692 | -89.6475
8.87 | -89.6527
8.8707 | -89.6559
8.8715 | -89.6603
8.8723 | -89.6639
8.8731 | -89.6656
8.8738 | -89.668
8.8746 | -89.6716
8.8754 | -89.6744
8.8762 | -89.6756
8.8769 | -89.6781
8.8777 | -89.6819
8.8785 | -89.6855
8.8793 | -89.6889
8.88 | -89.6927
8.8808 | -89.6979
8.8816 | -89.7022
8.8824 | -89.707
8.8831 | -89.7103
8.8839 | -89.7131
8.8847 | -89.716
8.8855 | -89.7193
8.8862 | -89.7226
8.887 | -89.7253
8.8878 | -89.7285
8.8886 | -89.7324
8.8893 | -89.7372
8.8901 | -89.7408
8.8909 | -89.7442
8.8917 | -89.7485
8.8924 | -89.7528
8.8932 | -89.7572
8.894 | -89.7606
8.8948 | -89.7631
8.8955 | -89.766
8.8963 | -89.77
8.8971 | -89.7736
8.8979 | -89.7763
8.8986 | -89.7782
8.8994 | -89.7816
8.9002 | -89.7852
8.901 | -89.7878
8.9017 | -89.7918
8.9025 | -89.7969
8.9033 | -89.8018
8.9041 | -89.8067
8.9048 | -89.8112
8.9056 | -89.8146
8.9064 | -89.8175
8.9072 | -89.8213
8.9079 | -89.8255
8.9087 | -89.8291
8.9095 | -89.833
8.9103 | -89.838
8.911 | -89.8433
8.9118 | -89.8464
8.9126 | -89.8487
8.9134 | -89.8514
8.9141 | -89.8542
8.9149 | -89.8573
8.9157 | -89.8616
8.9165 | -89.8661
8.9172 | -89.87
8.918 | -89.8752
8.9188 | -89.878
8.9196 | -89.8815
8.9203 | -89.8827
8.9211 | -89.8855
8.9219 | -89.8907
8.9227 | -89.8945
8.9234 | -89.8993
8.9242 | -89.9025
8.925 | -89.9054
8.9257 | -89.9082
8.9265 | -89.9114
8.9273 | -89.9152
8.9281 | -89.9195
8.9288 | -89.9233
8.9296 | -89.9258
8.9304 | -89.9287
8.9312 | -89.9325
8.9319 | -89.935
8.9327 | -89.9396
8.9335 | -89.9482
8.9343 | -89.9533
8.935 | -89.9569
8.9358 | -89.9605
8.9366 | -89.9635
8.9374 | -89.9675
8.9381 | -89.9722
8.9389 | -89.9762
8.9397 | -89.9794
8.9405 | -89.9825
8.9412 | -89.9852
8.942 | -89.9882
8.9428 | -89.9919
8.9436 | -89.995
8.9443 | -89.9984
8.9451 | -90.0014
8.9459 | -90.0052
8.9467 | -90.0087
8.9474 | -90.0117
8.9482 | -90.0144
8.949 | -90.02
8.9498 | -90.0239
8.9505 | -90.0286
8.9513 | -90.0377
8.9521 | -90.0419
8.9529 | -90.0465
8.9536 | -90.0523
8.9544 | -90.0557
8.9552 | -90.0582
8.956 | -90.0604
8.9567 | -90.0639
8.9575 | -90.0679
8.9583 | -90.072
8.9591 | -90.0771
8.9598 | -90.0812
8.9606 | -90.0838
8.9614 | -90.0876
8.9622 | -90.0908
8.9629 | -90.0947
8.9637 | -90.1015
8.9645 | -90.1081
8.9653 | -90.1103
8.966 | -90.1134
8.9668 | -90.1165
8.9676 | -90.1205
8.9684 | -90.1253
8.9691 | -90.1285
8.9699 | -90.1314
8.9707 | -90.1331
8.9715 | -90.1343
8.9722 | -90.1377
8.973 | -90.1404
8.9738 | -90.1415
8.9746 | -90.1436
8.9753 | -90.1479
8.9761 | -90.1513
8.9769 | -90.1543
8.9777 | -90.1572
8.9784 | -90.161
8.9792 | -90.1653
8.98 | -90.1681
8.9808 | -90.1708
8.9815 | -90.1735
8.9823 | -90.1763
8.9831 | -90.1795
8.9839 | -90.182
8.9846 | -90.1853
8.9854 | -90.1891
8.9862 | -90.192
8.987 | -90.1952
8.9877 | -90.1974
8.9885 | -90.2003
8.9893 | -90.2031
8.9901 | -90.2051
8.9908 | -90.2086
8.9916 | -90.2116
8.9924 | -90.215
8.9932 | -90.2191
8.9939 | -90.2222
8.9947 | -90.2252
8.9955 | -90.2276
8.9962 | -90.2302
8.997 | -90.2337
8.9978 | -90.237
8.9986 | -90.2404
8.9993 | -90.2431
9.0001 | -90.2483
9.0009 | -90.2534
9.0017 | -90.2584
9.0024 | -90.2625
9.0032 | -90.2656
9.004 | -90.2688
9.0048 | -90.2719
9.0055 | -90.2748
9.0063 | -90.2782
9.0071 | -90.2811
9.0079 | -90.2846
9.0086 | -90.2886
9.0094 | -90.2932
9.0102 | -90.2952
9.011 | -90.2985
9.0117 | -90.3005
9.0125 | -90.3038
9.0133 | -90.3069
9.0141 | -90.31
9.0148 | -90.3127
9.0156 | -90.3158
9.0164 | -90.3185
9.0172 | -90.3215
9.0179 | -90.3248
9.0187 | -90.3283
9.0195 | -90.3321
9.0203 | -90.3356
9.021 | -90.3393
9.0218 | -90.3422
9.0226 | -90.3446
9.0234 | -90.3477
9.0241 | -90.3511
9.0249 | -90.3542
9.0257 | -90.3572
9.0265 | -90.3584
9.0272 | -90.3611
9.028 | -90.3646
9.0288 | -90.3673
9.0296 | -90.3708
9.0303 | -90.3745
9.0311 | -90.3774
9.0319 | -90.3813
9.0327 | -90.3849
9.0334 | -90.3879
9.0342 | -90.3904
9.035 | -90.3923
9.0358 | -90.3946
9.0365 | -90.3968
9.0373 | -90.3996
9.0381 | -90.4026
9.0389 | -90.4053
9.0396 | -90.409
9.0404 | -90.4127
9.0412 | -90.4157
9.042 | -90.4178
9.0427 | -90.4203
9.0435 | -90.4235
9.0443 | -90.4257
9.0451 | -90.427
9.0458 | -90.4315
9.0466 | -90.4355
9.0474 | -90.4379
9.0482 | -90.441
9.0489 | -90.4439
9.0497 | -90.4465
9.0505 | -90.4512
9.0513 | -90.4566
9.052 | -90.4605
9.0528 | -90.4622
9.0536 | -90.4643
9.0544 | -90.4686
9.0551 | -90.4715
9.0559 | -90.4739
9.0567 | -90.4781
9.0575 | -90.4826
9.0582 | -90.4851
9.059 | -90.4871
9.0598 | -90.4899
9.0606 | -90.4923
9.0613 | -90.4944
9.0621 | -90.4972
9.0629 | -90.5002
9.0637 | -90.5045
9.0644 | -90.5084
9.0652 | -90.5106
9.066 | -90.5134
9.0668 | -90.5156
9.0675 | -90.5171
9.0683 | -90.5198
9.0691 | -90.5228
9.0699 | -90.5258
9.0706 | -90.5289
9.0714 | -90.5317
9.0722 | -90.5364
9.0729 | -90.5397
9.0737 | -90.5426
9.0745 | -90.5462
9.0753 | -90.5486
9.076 | -90.5534
9.0768 | -90.5594
9.0776 | -90.563
9.0784 | -90.5658
9.0791 | -90.5681
9.0799 | -90.5715
9.0807 | -90.5757
9.0815 | -90.5788
9.0822 | -90.5821
9.083 | -90.5853
9.0838 | -90.5864
9.0846 | -90.5881
9.0853 | -90.5927
9.0861 | -90.5965
9.0869 | -90.5985
9.0877 | -90.5999
9.0884 | -90.6024
9.0892 | -90.6053
9.09 | -90.6083
9.0908 | -90.6117
9.0915 | -90.6147
9.0923 | -90.6194
9.0931 | -90.6244
9.0939 | -90.627
9.0946 | -90.6332
9.0954 | -90.6404
9.0962 | -90.6428
9.097 | -90.6449
9.0977 | -90.6467
9.0985 | -90.6488
9.0993 | -90.6502
9.1001 | -90.652
9.1008 | -90.6557
9.1016 | -90.6586
9.1024 | -90.6613
9.1032 | -90.6641
9.1039 | -90.6662
9.1047 | -90.668
9.1055 | -90.6706
9.1063 | -90.6731
9.107 | -90.6755
9.1078 | -90.6775
9.1086 | -90.6803
9.1094 | -90.6845
9.1101 | -90.6882
9.1109 | -90.691
9.1117 | -90.6941
9.1125 | -90.6963
9.1132 | -90.6987
9.114 | -90.7009
9.1148 | -90.7026
9.1156 | -90.7075
9.1163 | -90.7103
9.1171 | -90.7131
9.1179 | -90.7163
9.1187 | -90.7189
9.1194 | -90.7215
9.1202 | -90.7237
9.121 | -90.7269
9.1218 | -90.7355
9.1225 | -90.74
9.1233 | -90.7424
9.1241 | -90.7446
9.1249 | -90.7484
9.1256 | -90.7519
9.1264 | -90.7554
9.1272 | -90.7574
9.128 | -90.7603
9.1287 | -90.7627
9.1295 | -90.7656
9.1303 | -90.7681
9.1311 | -90.7697
9.1318 | -90.7729
9.1326 | -90.7765
9.1334 | -90.7789
9.1342 | -90.7789
9.1349 | -90.7808
9.1357 | -90.784
9.1365 | -90.7871
9.1373 | -90.7901
9.138 | -90.7915
9.1388 | -90.793
9.1396 | -90.7952
9.1404 | -90.7979
9.1411 | -90.8011
9.1419 | -90.8039
9.1427 | -90.8062
9.1434 | -90.8084
9.1442 | -90.8109
9.145 | -90.8133
9.1458 | -90.8162
9.1465 | -90.8185
9.1473 | -90.8213
9.1481 | -90.8241
9.1489 | -90.8268
9.1496 | -90.8306
9.1504 | -90.8351
9.1512 | -90.8369
9.152 | -90.8378
9.1527 | -90.8404
9.1535 | -90.8423
9.1543 | -90.8441
9.1551 | -90.8469
9.1558 | -90.8512
9.1566 | -90.8532
9.1574 | -90.8555
9.1582 | -90.8579
9.1589 | -90.8605
9.1597 | -90.8644
9.1605 | -90.8691
9.1613 | -90.8729
9.162 | -90.8759
9.1628 | -90.878
9.1636 | -90.8803
9.1644 | -90.8831
9.1651 | -90.8857
9.1659 | -90.8887
9.1667 | -90.8901
9.1675 | -90.8922
9.1682 | -90.895
9.169 | -90.8976
9.1698 | -90.9002
9.1706 | -90.9038
9.1713 | -90.9043
9.1721 | -90.9036
9.1729 | -90.9054
9.1737 | -90.9073
9.1744 | -90.9082
9.1752 | -90.9095
9.176 | -90.9113
9.1768 | -90.9123
9.1775 | -90.9149
9.1783 | -90.9182
9.1791 | -90.9213
9.1799 | -90.9235
9.1806 | -90.9257
9.1814 | -90.9289
9.1822 | -90.9308
9.183 | -90.9318
9.1837 | -90.9339
9.1845 | -90.9351
9.1853 | -90.9375
9.1861 | -90.9403
9.1868 | -90.9405
9.1876 | -90.9412
9.1884 | -90.9429
9.1892 | -90.9451
9.1899 | -90.9462
9.1907 | -90.9467
9.1915 | -90.9485
9.1923 | -90.9507
9.193 | -90.9517
9.1938 | -90.952
9.1946 | -90.9523
9.1954 | -90.9537
9.1961 | -90.9552
9.1969 | -90.9565
9.1977 | -90.9574
9.1985 | -90.9583
9.1992 | -90.9576
9.2 | -90.957
9.2008 | -90.9564
9.2016 | -90.9573
9.2023 | -90.9593
9.2031 | -90.9611
9.2039 | -90.9632
9.2047 | -90.9648
9.2054 | -90.9664
9.2062 | -90.9681
9.207 | -90.97
9.2078 | -90.9715
9.2085 | -90.9732
9.2093 | -90.9752
9.2101 | -90.9774
9.2109 | -90.9799
9.2116 | -90.9812
9.2124 | -90.983
9.2132 | -90.985
9.214 | -90.9865
9.2147 | -90.9881
9.2155 | -90.9895
9.2163 | -90.9909
9.2171 | -90.9925
9.2178 | -90.9939
9.2186 | -90.9939
9.2194 | -90.9958
9.2201 | -90.9988
9.2209 | -91.0004
9.2217 | -91.0022
9.2225 | -91.004
9.2232 | -91.0043
9.224 | -91.0081
9.2248 | -91.0108
9.2256 | -91.0133
9.2263 | -91.0157
9.2271 | -91.0165
9.2279 | -91.0167
9.2287 | -91.0173
9.2294 | -91.0185
9.2302 | -91.0198
9.231 | -91.0211
9.2318 | -91.0224
9.2325 | -91.0237
9.2333 | -91.0248
9.2341 | -91.0259
9.2349 | -91.0274
9.2356 | -91.0281
9.2364 | -91.028
9.2372 | -91.0272
9.238 | -91.0275
9.2387 | -91.0285
9.2395 | -91.0296
9.2403 | -91.0307
9.2411 | -91.0324
9.2418 | -91.0329
9.2426 | -91.0332
9.2434 | -91.0352
9.2442 | -91.0362
9.2449 | -91.0387
9.2457 | -91.0422
9.2465 | -91.0427
9.2473 | -91.0433
9.248 | -91.0436
9.2488 | -91.0446
9.2496 | -91.0468
9.2504 | -91.0497
9.2511 | -91.0504
9.2519 | -91.0502
9.2527 | -91.0505
9.2535 | -91.0503
9.2542 | -91.0528
9.255 | -91.0537
9.2558 | -91.0545
9.2566 | -91.056
9.2573 | -91.0568
9.2581 | -91.0581
9.2589 | -91.06
9.2597 | -91.0599
9.2604 | -91.0583
9.2612 | -91.0576
9.262 | -91.0592
9.2628 | -91.0605
9.2635 | -91.0616
9.2643 | -91.0615
9.2651 | -91.0596
9.2659 | -91.0607
9.2666 | -91.0614
9.2674 | -91.0617
9.2682 | -91.0621
9.269 | -91.0625
9.2697 | -91.0634
9.2705 | -91.0644
9.2713 | -91.0652
9.2721 | -91.0661
9.2728 | -91.0665
9.2736 | -91.0673
9.2744 | -91.0677
9.2752 | -91.0705
9.2759 | -91.0704
9.2767 | -91.0723
9.2775 | -91.0731
9.2783 | -91.0741
9.279 | -91.0758
9.2798 | -91.0763
9.2806 | -91.0774
9.2814 | -91.0785
9.2821 | -91.0788
9.2829 | -91.08
9.2837 | -91.0799
9.2845 | -91.0805
9.2852 | -91.0802
9.286 | -91.0799
9.2868 | -91.0801
9.2876 | -91.0805
9.2883 | -91.0805
9.2891 | -91.0813
9.2899 | -91.0814
9.2906 | -91.0814
9.2914 | -91.0818
9.2922 | -91.0829
9.293 | -91.0837
9.2937 | -91.0849
9.2945 | -91.086
9.2953 | -91.0864
9.2961 | -91.0872
9.2968 | -91.0856
9.2976 | -91.0838
9.2984 | -91.0829
9.2992 | -91.0832
9.2999 | -91.0828
9.3007 | -91.0836
9.3015 | -91.0855
9.3023 | -91.0878
9.303 | -91.088
9.3038 | -91.0881
9.3046 | -91.0878
9.3054 | -91.0849
9.3061 | -91.0847
9.3069 | -91.0834
9.3077 | -91.0823
9.3085 | -91.0819
9.3092 | -91.0813
9.31 | -91.0811
9.3108 | -91.08
9.3116 | -91.079
9.3123 | -91.0818
9.3131 | -91.0829
9.3139 | -91.0803
9.3147 | -91.079
9.3154 | -91.0787
9.3162 | -91.0783
9.317 | -91.0779
9.3178 | -91.08
9.3185 | -91.0805
9.3193 | -91.0813
9.3201 | -91.0829
9.3209 | -91.0827
9.3216 | -91.0807
9.3224 | -91.0789
9.3232 | -91.078
9.324 | -91.078
9.3247 | -91.0769
9.3255 | -91.0758
9.3263 | -91.0744
9.3271 | -91.0734
9.3278 | -91.0729
9.3286 | -91.0725
9.3294 | -91.0714
9.3302 | -91.0711
9.3309 | -91.0737
9.3317 | -91.0748
9.3325 | -91.0739
9.3333 | -91.0735
9.334 | -91.0722
9.3348 | -91.0701
9.3356 | -91.0697
9.3364 | -91.0694
9.3371 | -91.0681
9.3379 | -91.0676
9.3387 | -91.0672
9.3395 | -91.0666
9.3402 | -91.0668
9.341 | -91.0664
9.3418 | -91.0658
9.3426 | -91.0649
9.3433 | -91.0624
9.3441 | -91.0618
9.3449 | -91.06
9.3457 | -91.0593
9.3464 | -91.059
9.3472 | -91.0588
9.348 | -91.0588
9.3488 | -91.0593
9.3495 | -91.0619
9.3503 | -91.0621
9.3511 | -91.0614
9.3519 | -91.0614
9.3526 | -91.0612
9.3534 | -91.0615
9.3542 | -91.0612
9.355 | -91.0601
9.3557 | -91.0581
9.3565 | -91.057
9.3573 | -91.0563
9.3581 | -91.0549
9.3588 | -91.0532
9.3596 | -91.0523
9.3604 | -91.0515
9.3612 | -91.0508
9.3619 | -91.0497
9.3627 | -91.0464
9.3635 | -91.0453
9.3643 | -91.0438
9.365 | -91.0411
9.3658 | -91.0366
9.3666 | -91.0356
9.3673 | -91.0346
9.3681 | -91.0336
9.3689 | -91.0317
9.3697 | -91.0304
9.3704 | -91.0299
9.3712 | -91.0298
9.372 | -91.026
9.3728 | -91.025
9.3735 | -91.0238
9.3743 | -91.0226
9.3751 | -91.0214
9.3759 | -91.0202
9.3766 | -91.019
9.3774 | -91.0186
9.3782 | -91.0186
9.379 | -91.0178
9.3797 | -91.0177
9.3805 | -91.0173
9.3813 | -91.0164
9.3821 | -91.0163
9.3828 | -91.0152
9.3836 | -91.0137
9.3844 | -91.0117
9.3852 | -91.0083
9.3859 | -91.003
9.3867 | -91.0001
9.3875 | -91.0001
9.3883 | -90.999
9.389 | -90.9962
9.3898 | -90.9944
9.3906 | -90.9931
9.3914 | -90.9919
9.3921 | -90.9903
9.3929 | -90.9875
9.3937 | -90.9856
9.3945 | -90.9838
9.3952 | -90.9827
9.396 | -90.9805
9.3968 | -90.9774
9.3976 | -90.9747
9.3983 | -90.9732
9.3991 | -90.9703
9.3999 | -90.9688
9.4007 | -90.9674
9.4014 | -90.9662
9.4022 | -90.9627
9.403 | -90.9627
9.4038 | -90.9619
9.4045 | -90.9597
9.4053 | -90.9569
9.4061 | -90.9548
9.4069 | -90.9525
9.4076 | -90.9497
9.4084 | -90.9489
9.4092 | -90.9492
9.41 | -90.9476
9.4107 | -90.9465
9.4115 | -90.9459
9.4123 | -90.9464
9.4131 | -90.9461
9.4138 | -90.9442
9.4146 | -90.9423
9.4154 | -90.9398
9.4162 | -90.9377
9.4169 | -90.9369
9.4177 | -90.9349
9.4185 | -90.933
9.4193 | -90.9309
9.42 | -90.9286
9.4208 | -90.928
9.4216 | -90.9253
9.4224 | -90.9229
9.4231 | -90.9211
9.4239 | -90.9189
9.4247 | -90.9162
9.4255 | -90.9127
9.4262 | -90.9078
9.427 | -90.9043
9.4278 | -90.9021
9.4286 | -90.9014
9.4293 | -90.9016
9.4301 | -90.9008
9.4309 | -90.8983
9.4317 | -90.8964
9.4324 | -90.8925
9.4332 | -90.8905
9.434 | -90.8887
9.4348 | -90.8862
9.4355 | -90.8838
9.4363 | -90.8818
9.4371 | -90.8797
9.4379 | -90.8773
9.4386 | -90.8748
9.4394 | -90.873
9.4402 | -90.8708
9.441 | -90.869
9.4417 | -90.8678
9.4425 | -90.8669
9.4433 | -90.8662
9.444 | -90.8642
9.4448 | -90.8624
9.4456 | -90.86
9.4464 | -90.858
9.4471 | -90.8555
9.4479 | -90.8524
9.4487 | -90.85
9.4495 | -90.8479
9.4502 | -90.846
9.451 | -90.8438
9.4518 | -90.8417
9.4526 | -90.839
9.4533 | -90.8357
9.4541 | -90.8326
9.4549 | -90.8309
9.4557 | -90.8302
9.4564 | -90.8274
9.4572 | -90.8246
9.458 | -90.8222
9.4588 | -90.8199
9.4595 | -90.8176
9.4603 | -90.8145
9.4611 | -90.8115
9.4619 | -90.8086
9.4626 | -90.8061
9.4634 | -90.8044
9.4642 | -90.8021
9.465 | -90.7995
9.4657 | -90.7973
9.4665 | -90.7946
9.4673 | -90.7919
9.4681 | -90.7901
9.4688 | -90.7873
9.4696 | -90.785
9.4704 | -90.7832
9.4712 | -90.781
9.4719 | -90.7783
9.4727 | -90.7762
9.4735 | -90.7739
9.4743 | -90.7713
9.475 | -90.7664
9.4758 | -90.7605
9.4766 | -90.7564
9.4774 | -90.753
9.4781 | -90.7499
9.4789 | -90.7463
9.4797 | -90.7426
9.4805 | -90.7399
9.4812 | -90.7372
9.482 | -90.7349
9.4828 | -90.7326
9.4836 | -90.7304
9.4843 | -90.7283
9.4851 | -90.726
9.4859 | -90.7237
9.4867 | -90.7219
9.4874 | -90.7197
9.4882 | -90.7182
9.489 | -90.7169
9.4898 | -90.7148
9.4905 | -90.7126
9.4913 | -90.7103
9.4921 | -90.7079
9.4929 | -90.705
9.4936 | -90.701
9.4944 | -90.6974
9.4952 | -90.6936
9.496 | -90.6905
9.4967 | -90.6882
9.4975 | -90.6874
9.4983 | -90.6847
9.4991 | -90.6812
9.4998 | -90.6755
9.5006 | -90.6731
9.5014 | -90.6705
9.5022 | -90.6679
9.5029 | -90.6663
9.5037 | -90.6646
9.5045 | -90.6608
9.5053 | -90.6556
9.506 | -90.6526
9.5068 | -90.6498
9.5076 | -90.6472
9.5084 | -90.6449
9.5091 | -90.642
9.5099 | -90.6397
9.5107 | -90.6379
9.5115 | -90.6369
9.5122 | -90.6351
9.513 | -90.6301
9.5138 | -90.6258
9.5145 | -90.6242
9.5153 | -90.6237
9.5161 | -90.6212
9.5169 | -90.6182
9.5176 | -90.6152
9.5184 | -90.612
9.5192 | -90.6093
9.52 | -90.6062
9.5207 | -90.6034
9.5215 | -90.6009
9.5223 | -90.5985
9.5231 | -90.5965
9.5238 | -90.5945
9.5246 | -90.5911
9.5254 | -90.5879
9.5262 | -90.5858
9.5269 | -90.5834
9.5277 | -90.5793
9.5285 | -90.576
9.5293 | -90.5729
9.53 | -90.5692
9.5308 | -90.5658
9.5316 | -90.5634
9.5324 | -90.561
9.5331 | -90.5589
9.5339 | -90.5564
9.5347 | -90.5538
9.5355 | -90.5506
9.5362 | -90.548
9.537 | -90.5452
9.5378 | -90.5426
9.5386 | -90.5412
9.5393 | -90.5399
9.5401 | -90.5369
9.5409 | -90.5335
9.5417 | -90.5315
9.5424 | -90.5297
9.5432 | -90.5272
9.544 | -90.5238
9.5448 | -90.5204
9.5455 | -90.5175
9.5463 | -90.515
9.5471 | -90.513
9.5479 | -90.5107
9.5486 | -90.5079
9.5494 | -90.505
9.5502 | -90.5026
9.551 | -90.5005
9.5517 | -90.4983
9.5525 | -90.4961
9.5533 | -90.4939
9.5541 | -90.4919
9.5548 | -90.4898
9.5556 | -90.4879
9.5564 | -90.4854
9.5572 | -90.481
9.5579 | -90.4769
9.5587 | -90.4737
9.5595 | -90.4705
9.5603 | -90.4673
9.561 | -90.4652
9.5618 | -90.4629
9.5626 | -90.4621
9.5634 | -90.4607
9.5641 | -90.4579
9.5649 | -90.4544
9.5657 | -90.4503
9.5665 | -90.4462
9.5672 | -90.4422
9.568 | -90.4395
9.5688 | -90.4372
9.5696 | -90.4346
9.5703 | -90.4321
9.5711 | -90.4292
9.5719 | -90.4267
9.5727 | -90.4246
9.5734 | -90.4226
9.5742 | -90.4203
9.575 | -90.4191
9.5758 | -90.4172
9.5765 | -90.4148
9.5773 | -90.4121
9.5781 | -90.4086
9.5789 | -90.4057
9.5796 | -90.4036
9.5804 | -90.4017
9.5812 | -90.3997
9.582 | -90.397
9.5827 | -90.3943
9.5835 | -90.392
9.5843 | -90.39
9.5851 | -90.3871
9.5858 | -90.3842
9.5866 | -90.3819
9.5874 | -90.3794
9.5882 | -90.3771
9.5889 | -90.3754
9.5897 | -90.373
9.5905 | -90.37
9.5912 | -90.3651
9.592 | -90.3614
9.5928 | -90.3592
9.5936 | -90.3567
9.5943 | -90.3533
9.5951 | -90.3508
9.5959 | -90.3484
9.5967 | -90.3461
9.5974 | -90.3439
9.5982 | -90.3419
9.599 | -90.3396
9.5998 | -90.3366
9.6005 | -90.334
9.6013 | -90.3303
9.6021 | -90.3279
9.6029 | -90.3257
9.6036 | -90.3235
9.6044 | -90.3213
9.6052 | -90.3192
9.606 | -90.3171
9.6067 | -90.3153
9.6075 | -90.3114
9.6083 | -90.3088
9.6091 | -90.3071
9.6098 | -90.3052
9.6106 | -90.3027
9.6114 | -90.3004
9.6122 | -90.298
9.6129 | -90.2959
9.6137 | -90.2938
9.6145 | -90.2914
9.6153 | -90.288
9.616 | -90.2869
9.6168 | -90.2858
9.6176 | -90.2858
9.6184 | -90.2846
9.6191 | -90.283
9.6199 | -90.2817
9.6207 | -90.2802
9.6215 | -90.2778
9.6222 | -90.2758
9.623 | -90.2736
9.6238 | -90.272
9.6246 | -90.2704
9.6253 | -90.2682
9.6261 | -90.2656
9.6269 | -90.2635
9.6277 | -90.2614
9.6284 | -90.2591
9.6292 | -90.2566
9.63 | -90.2547
9.6308 | -90.2523
9.6315 | -90.2496
9.6323 | -90.2475
9.6331 | -90.2454
9.6339 | -90.2445
9.6346 | -90.2425
9.6354 | -90.2413
9.6362 | -90.2397
9.637 | -90.2381
9.6377 | -90.2379
9.6385 | -90.2367
9.6393 | -90.2339
9.6401 | -90.2293
9.6408 | -90.2261
9.6416 | -90.2234
9.6424 | -90.2211
9.6432 | -90.2187
9.6439 | -90.2171
9.6447 | -90.2159
9.6455 | -90.2138
9.6463 | -90.2119
9.647 | -90.2103
9.6478 | -90.2084
9.6486 | -90.2064
9.6494 | -90.2048
9.6501 | -90.203
9.6509 | -90.2014
9.6517 | -90.1983
9.6525 | -90.1964
9.6532 | -90.1947
9.654 | -90.1926
9.6548 | -90.19
9.6556 | -90.1877
9.6563 | -90.1872
9.6571 | -90.1846
9.6579 | -90.1825
9.6587 | -90.1816
9.6594 | -90.1805
9.6602 | -90.179
9.661 | -90.1785
9.6617 | -90.1771
9.6625 | -90.174
9.6633 | -90.1712
9.6641 | -90.1686
9.6648 | -90.1663
9.6656 | -90.1649
9.6664 | -90.1631
9.6672 | -90.161
9.6679 | -90.1591
9.6687 | -90.1566
9.6695 | -90.1548
9.6703 | -90.1534
9.671 | -90.1525
9.6718 | -90.1516
9.6726 | -90.1501
9.6734 | -90.1485
9.6741 | -90.1471
9.6749 | -90.1444
9.6757 | -90.1434
9.6765 | -90.1431
9.6772 | -90.1419
9.678 | -90.1399
9.6788 | -90.1385
9.6796 | -90.1371
9.6803 | -90.1358
9.6811 | -90.1345
9.6819 | -90.1325
9.6827 | -90.1299
9.6834 | -90.1284
9.6842 | -90.1278
9.685 | -90.1272
9.6858 | -90.127
9.6865 | -90.1258
9.6873 | -90.1242
9.6881 | -90.1208
9.6889 | -90.1192
9.6896 | -90.1189
9.6904 | -90.1176
9.6912 | -90.116
9.692 | -90.1148
9.6927 | -90.1136
9.6935 | -90.1118
9.6943 | -90.1082
9.6951 | -90.1065
9.6958 | -90.1064
9.6966 | -90.1047
9.6974 | -90.1032
9.6982 | -90.1021
9.6989 | -90.0992
9.6997 | -90.0961
9.7005 | -90.0947
9.7013 | -90.094
9.702 | -90.0928
9.7028 | -90.0907
9.7036 | -90.0895
9.7044 | -90.0884
9.7051 | -90.0884
9.7059 | -90.0877
9.7067 | -90.0865
9.7075 | -90.0858
9.7082 | -90.0847
9.709 | -90.0836
9.7098 | -90.0821
9.7106 | -90.0813
9.7113 | -90.0802
9.7121 | -90.079
9.7129 | -90.078
9.7137 | -90.0773
9.7144 | -90.076
9.7152 | -90.0746
9.716 | -90.0734
9.7168 | -90.0721
9.7175 | -90.0714
9.7183 | -90.0707
9.7191 | -90.0705
9.7199 | -90.0706
9.7206 | -90.0697
9.7214 | -90.0688
9.7222 | -90.0681
9.723 | -90.0685
9.7237 | -90.0679
9.7245 | -90.0674
9.7253 | -90.0665
9.7261 | -90.0655
9.7268 | -90.0647
9.7276 | -90.0639
9.7284 | -90.0635
9.7292 | -90.0633
9.7299 | -90.0625
9.7307 | -90.0616
9.7315 | -90.0609
9.7323 | -90.0601
9.733 | -90.059
9.7338 | -90.0578
9.7346 | -90.0574
9.7354 | -90.0568
9.7361 | -90.0563
9.7369 | -90.0561
9.7377 | -90.0561
9.7384 | -90.0564
9.7392 | -90.055
9.74 | -90.0539
9.7408 | -90.0537
9.7415 | -90.0534
9.7423 | -90.0523
9.7431 | -90.0515
9.7439 | -90.0505
9.7446 | -90.0493
9.7454 | -90.0478
9.7462 | -90.0458
9.747 | -90.0454
9.7477 | -90.0449
9.7485 | -90.0442
9.7493 | -90.0434
9.7501 | -90.0425
9.7508 | -90.0426
9.7516 | -90.0429
9.7524 | -90.0426
9.7532 | -90.0422
9.7539 | -90.0417
9.7547 | -90.0422
9.7555 | -90.0426
9.7563 | -90.0419
9.757 | -90.0416
9.7578 | -90.041
9.7586 | -90.0405
9.7594 | -90.0404
9.7601 | -90.0405
9.7609 | -90.0403
9.7617 | -90.04
9.7625 | -90.0399
9.7632 | -90.0401
9.764 | -90.0399
9.7648 | -90.0397
9.7656 | -90.0397
9.7663 | -90.0394
9.7671 | -90.0383
9.7679 | -90.0394
9.7687 | -90.0398
9.7694 | -90.04
9.7702 | -90.0395
9.771 | -90.0396
9.7718 | -90.0396
9.7725 | -90.0397
9.7733 | -90.0396
9.7741 | -90.0396
9.7749 | -90.0398
9.7756 | -90.04
9.7764 | -90.039
9.7772 | -90.0394
9.778 | -90.0392
9.7787 | -90.0385
9.7795 | -90.0376
9.7803 | -90.0368
9.7811 | -90.0372
9.7818 | -90.0373
9.7826 | -90.0377
9.7834 | -90.0375
9.7842 | -90.0381
9.7849 | -90.0382
9.7857 | -90.0378
9.7865 | -90.0377
9.7873 | -90.0383
9.788 | -90.0388
9.7888 | -90.0389
9.7896 | -90.0391
9.7904 | -90.0397
9.7911 | -90.0395
9.7919 | -90.0401
9.7927 | -90.041
9.7935 | -90.0412
9.7942 | -90.0414
9.795 | -90.0413
9.7958 | -90.0421
9.7966 | -90.0428
9.7973 | -90.0436
9.7981 | -90.0429
9.7989 | -90.0422
9.7997 | -90.0422
9.8004 | -90.0427
9.8012 | -90.0434
9.802 | -90.0429
9.8028 | -90.0416
9.8035 | -90.0419
9.8043 | -90.0413
9.8051 | -90.0394
9.8059 | -90.0388
9.8066 | -90.0391
9.8074 | -90.0378
9.8082 | -90.0373
9.8089 | -90.037
9.8097 | -90.0372
9.8105 | -90.0382
9.8113 | -90.0387
9.812 | -90.0404
9.8128 | -90.0418
9.8136 | -90.0439
9.8144 | -90.0452
9.8151 | -90.0446
9.8159 | -90.0441
9.8167 | -90.0431
9.8175 | -90.0426
9.8182 | -90.0429
9.819 | -90.0431
9.8198 | -90.0438
9.8206 | -90.043
9.8213 | -90.0429
9.8221 | -90.0434
9.8229 | -90.0439
9.8237 | -90.044
9.8244 | -90.0441
9.8252 | -90.0443
9.826 | -90.0444
9.8268 | -90.0452
9.8275 | -90.046
9.8283 | -90.0471
9.8291 | -90.0478
9.8299 | -90.0483
9.8306 | -90.049
9.8314 | -90.0513
9.8322 | -90.0533
9.833 | -90.0532
9.8337 | -90.0527
9.8345 | -90.0539
9.8353 | -90.0559
9.8361 | -90.0561
9.8368 | -90.0571
9.8376 | -90.058
9.8384 | -90.0587
9.8392 | -90.0596
9.8399 | -90.06
9.8407 | -90.0598
9.8415 | -90.0598
9.8423 | -90.06
9.843 | -90.0611
9.8438 | -90.0623
9.8446 | -90.0632
9.8454 | -90.0649
9.8461 | -90.0658
9.8469 | -90.0669
9.8477 | -90.0679
9.8485 | -90.0702
9.8492 | -90.0721
9.85 | -90.0723
9.8508 | -90.0723
9.8516 | -90.0739
9.8523 | -90.0758
9.8531 | -90.0767
9.8539 | -90.0774
9.8547 | -90.0778
9.8554 | -90.079
9.8562 | -90.0807
9.857 | -90.0824
9.8578 | -90.0828
9.8585 | -90.0837
9.8593 | -90.0849
9.8601 | -90.086
9.8609 | -90.0879
9.8616 | -90.0912
9.8624 | -90.0946
9.8632 | -90.0973
9.864 | -90.0992
9.8647 | -90.1
9.8655 | -90.1
9.8663 | -90.0997
9.8671 | -90.1
9.8678 | -90.1004
9.8686 | -90.1012
9.8694 | -90.1023
9.8702 | -90.1022
9.8709 | -90.1012
9.8717 | -90.1018
9.8725 | -90.1024
9.8733 | -90.1033
9.874 | -90.1043
9.8748 | -90.1055
9.8756 | -90.1066
9.8764 | -90.1074
9.8771 | -90.1083
9.8779 | -90.1088
9.8787 | -90.1082
9.8795 | -90.1097
9.8802 | -90.1107
9.881 | -90.1117
9.8818 | -90.113
9.8826 | -90.1146
9.8833 | -90.1148
9.8841 | -90.1166
9.8849 | -90.1188
9.8856 | -90.1202
9.8864 | -90.1208
9.8872 | -90.1226
9.888 | -90.124
9.8887 | -90.1249
9.8895 | -90.1259
9.8903 | -90.1272
9.8911 | -90.1293
9.8918 | -90.1311
9.8926 | -90.1326
9.8934 | -90.1336
9.8942 | -90.1344
9.8949 | -90.1348
9.8957 | -90.1346
9.8965 | -90.1337
9.8973 | -90.1336
9.898 | -90.1354
9.8988 | -90.1369
9.8996 | -90.1382
9.9004 | -90.1402
9.9011 | -90.1425
9.9019 | -90.1447
9.9027 | -90.1469
9.9035 | -90.149
9.9042 | -90.1502
9.905 | -90.1509
9.9058 | -90.1522
9.9066 | -90.1542
9.9073 | -90.1563
9.9081 | -90.159
9.9089 | -90.1613
9.9097 | -90.1631
9.9104 | -90.1645
9.9112 | -90.1658
9.912 | -90.1674
9.9128 | -90.1686
9.9135 | -90.1704
9.9143 | -90.1719
9.9151 | -90.1732
9.9159 | -90.1745
9.9166 | -90.1756
9.9174 | -90.1775
9.9182 | -90.1807
9.919 | -90.1818
9.9197 | -90.182
9.9205 | -90.1828
9.9213 | -90.184
9.9221 | -90.1852
9.9228 | -90.1863
9.9236 | -90.1881
9.9244 | -90.1903
9.9252 | -90.1919
9.9259 | -90.1935
9.9267 | -90.195
9.9275 | -90.1965
9.9283 | -90.1978
9.929 | -90.1988
9.9298 | -90.1992
9.9306 | -90.2005
9.9314 | -90.2023
9.9321 | -90.2057
9.9329 | -90.2073
9.9337 | -90.2084
9.9345 | -90.2092
9.9352 | -90.2103
9.936 | -90.2118
9.9368 | -90.213
9.9376 | -90.2145
9.9383 | -90.2161
9.9391 | -90.2173
9.9399 | -90.219
9.9407 | -90.2207
9.9414 | -90.2222
9.9422 | -90.2235
9.943 | -90.2256
9.9438 | -90.2271
9.9445 | -90.2281
9.9453 | -90.2271
9.9461 | -90.2262
9.9469 | -90.2273
9.9476 | -90.2288
9.9484 | -90.23
9.9492 | -90.2314
9.95 | -90.2333
9.9507 | -90.2347
9.9515 | -90.2357
9.9523 | -90.2367
9.9531 | -90.2381
9.9538 | -90.2402
9.9546 | -90.2423
9.9554 | -90.2444
9.9561 | -90.2462
9.9569 | -90.2479
9.9577 | -90.2498
9.9585 | -90.2511
9.9592 | -90.2524
9.96 | -90.2537
9.9608 | -90.2547
9.9616 | -90.2557
9.9623 | -90.2564
9.9631 | -90.2579
9.9639 | -90.2593
9.9647 | -90.2614
9.9654 | -90.2641
9.9662 | -90.2668
9.967 | -90.2689
9.9678 | -90.2694
9.9685 | -90.2705
9.9693 | -90.272
9.9701 | -90.2731
9.9709 | -90.2743
9.9716 | -90.2763
9.9724 | -90.2767
9.9732 | -90.2778
9.974 | -90.2795
9.9747 | -90.2813
9.9755 | -90.285
9.9763 | -90.2867
9.9771 | -90.2872
9.9778 | -90.2875
9.9786 | -90.2889
9.9794 | -90.2896
9.9802 | -90.2889
9.9809 | -90.2896
9.9817 | -90.2922
9.9825 | -90.2935
9.9833 | -90.2959
9.984 | -90.2997
9.9848 | -90.302
9.9856 | -90.3032
9.9864 | -90.304
9.9871 | -90.305
9.9879 | -90.3059
9.9887 | -90.3072
9.9895 | -90.3087
9.9902 | -90.3097
9.991 | -90.3106
9.9918 | -90.3113
9.9926 | -90.309
9.9933 | -90.3088
9.9941 | -90.3108
9.9949 | -90.3124
9.9957 | -90.3134
9.9964 | -90.3145
9.9972 | -90.316
9.998 | -90.3178
9.9988 | -90.3198
9.9995 | -90.3211
10.0003 | -90.3225
10.0011 | -90.3243
10.0019 | -90.3265
10.0026 | -90.3286
10.0034 | -90.3294
10.0042 | -90.3292
10.005 | -90.3308
10.0057 | -90.3331
10.0065 | -90.3354
10.0073 | -90.3369
10.0081 | -90.3381
10.0088 | -90.3399
10.0096 | -90.3417
10.0104 | -90.343
10.0112 | -90.3437
10.0119 | -90.3446
10.0127 | -90.3457
10.0135 | -90.3472
10.0143 | -90.3487
10.015 | -90.3501
10.0158 | -90.3522
10.0166 | -90.3543
10.0174 | -90.355
10.0181 | -90.357
10.0189 | -90.3588
10.0197 | -90.3599
10.0205 | -90.3608
10.0212 | -90.3614
10.022 | -90.3625
10.0228 | -90.3643
10.0236 | -90.3659
10.0243 | -90.3672
10.0251 | -90.3684
10.0259 | -90.3695
10.0267 | -90.371
10.0274 | -90.3726
10.0282 | -90.3744
10.029 | -90.3771
10.0297 | -90.379
10.0305 | -90.3797
10.0313 | -90.3811
10.0321 | -90.3828
10.0328 | -90.3842
10.0336 | -90.3858
10.0344 | -90.3874
10.0352 | -90.3892
10.0359 | -90.3894
10.0367 | -90.3923
10.0375 | -90.3926
10.0383 | -90.3935
10.039 | -90.3968
10.0398 | -90.3996
10.0406 | -90.4014
10.0414 | -90.4029
10.0421 | -90.4038
10.0429 | -90.4046
10.0437 | -90.406
10.0445 | -90.4065
10.0452 | -90.407
10.046 | -90.4076
10.0468 | -90.409
10.0476 | -90.4075
10.0483 | -90.4069
10.0491 | -90.4077
10.0499 | -90.4081
10.0507 | -90.4088
10.0514 | -90.4097
10.0522 | -90.4108
10.053 | -90.4114
10.0538 | -90.4119
10.0545 | -90.4126
10.0553 | -90.4125
10.0561 | -90.4133
10.0569 | -90.4139
10.0576 | -90.415
10.0584 | -90.4162
10.0592 | -90.4175
10.06 | -90.4187
10.0607 | -90.4206
10.0615 | -90.4249
10.0623 | -90.4277
10.0631 | -90.4288
10.0638 | -90.4296
10.0646 | -90.4303
10.0654 | -90.432
10.0662 | -90.4334
10.0669 | -90.4346
10.0677 | -90.4355
10.0685 | -90.4365
10.0693 | -90.4384
10.07 | -90.439
10.0708 | -90.4381
10.0716 | -90.4375
10.0724 | -90.4385
10.0731 | -90.4406
10.0739 | -90.4427
10.0747 | -90.4435
10.0755 | -90.4455
10.0762 | -90.446
10.077 | -90.4462
10.0778 | -90.449
10.0786 | -90.4512
10.0793 | -90.4522
10.0801 | -90.4529
10.0809 | -90.4538
10.0817 | -90.4542
10.0824 | -90.4553
10.0832 | -90.4563
10.084 | -90.4578
10.0848 | -90.4593
10.0855 | -90.4605
10.0863 | -90.4601
10.0871 | -90.4595
10.0879 | -90.4603
10.0886 | -90.4611
10.0894 | -90.4622
10.0902 | -90.4633
10.091 | -90.4638
10.0917 | -90.4633
10.0925 | -90.4624
10.0933 | -90.4637
10.0941 | -90.4654
10.0948 | -90.466
10.0956 | -90.4671
10.0964 | -90.4677
10.0972 | -90.4682
10.0979 | -90.4678
10.0987 | -90.4663
10.0995 | -90.4664
10.1003 | -90.4668
10.101 | -90.4668
10.1018 | -90.4675
10.1026 | -90.4691
10.1034 | -90.4703
10.1041 | -90.472
10.1049 | -90.475
10.1057 | -90.4768
10.1064 | -90.4774
10.1072 | -90.4777
10.108 | -90.4772
10.1088 | -90.4766
10.1095 | -90.477
10.1103 | -90.4772
10.1111 | -90.4781
10.1119 | -90.4795
10.1126 | -90.4803
10.1134 | -90.4803
10.1142 | -90.4805
10.115 | -90.4805
10.1157 | -90.4812
10.1165 | -90.4816
10.1173 | -90.4821
10.1181 | -90.4823
10.1188 | -90.4811
10.1196 | -90.481
10.1204 | -90.4813
10.1212 | -90.4818
10.1219 | -90.4819
10.1227 | -90.4819
10.1235 | -90.482
10.1243 | -90.4816
10.125 | -90.48
10.1258 | -90.4798
10.1266 | -90.4799
10.1274 | -90.4806
10.1281 | -90.4812
10.1289 | -90.4817
10.1297 | -90.4821
10.1305 | -90.4822
10.1312 | -90.4816
10.132 | -90.4806
10.1328 | -90.4806
10.1336 | -90.4822
10.1343 | -90.4836
10.1351 | -90.4839
10.1359 | -90.4837
10.1367 | -90.4836
10.1374 | -90.484
10.1382 | -90.4843
10.139 | -90.4847
10.1398 | -90.4849
10.1405 | -90.4845
10.1413 | -90.4855
10.1421 | -90.4863
10.1429 | -90.4869
10.1436 | -90.4875
10.1444 | -90.4871
10.1452 | -90.487
10.146 | -90.4877
10.1467 | -90.4872
10.1475 | -90.4872
10.1483 | -90.4869
10.1491 | -90.4874
10.1498 | -90.4875
10.1506 | -90.4879
10.1514 | -90.4879
10.1522 | -90.4861
10.1529 | -90.4845
10.1537 | -90.4845
10.1545 | -90.4846
10.1553 | -90.4849
10.156 | -90.4844
10.1568 | -90.4856
10.1576 | -90.4881
10.1584 | -90.4882
10.1591 | -90.4881
10.1599 | -90.4884
10.1607 | -90.4892
10.1615 | -90.4902
10.1622 | -90.4915
10.163 | -90.4916
10.1638 | -90.491
10.1646 | -90.4907
10.1653 | -90.4898
10.1661 | -90.4893
10.1669 | -90.489
10.1677 | -90.4882
10.1684 | -90.4872
10.1692 | -90.4868
10.17 | -90.4868
10.1708 | -90.4868
10.1715 | -90.4862
10.1723 | -90.4861
10.1731 | -90.4857
10.1739 | -90.4855
10.1746 | -90.4855
10.1754 | -90.4852
10.1762 | -90.4855
10.1769 | -90.4862
10.1777 | -90.4876
10.1785 | -90.4891
10.1793 | -90.4903
10.18 | -90.4908
10.1808 | -90.4895
10.1816 | -90.489
10.1824 | -90.4904
10.1831 | -90.4931
10.1839 | -90.4933
10.1847 | -90.4928
10.1855 | -90.4927
10.1862 | -90.4919
10.187 | -90.4917
10.1878 | -90.4912
10.1886 | -90.4905
10.1893 | -90.4921
10.1901 | -90.4917
10.1909 | -90.4908
10.1917 | -90.4893
10.1924 | -90.4878
10.1932 | -90.4873
10.194 | -90.4879
10.1948 | -90.4877
10.1955 | -90.4877
10.1963 | -90.4882
10.1971 | -90.4886
10.1979 | -90.4906
10.1986 | -90.4917
10.1994 | -90.4904
10.2002 | -90.4901
10.201 | -90.4907
10.2017 | -90.4906
10.2025 | -90.4904
10.2033 | -90.4907
10.2041 | -90.491
10.2048 | -90.4913
10.2056 | -90.4908
10.2064 | -90.4902
10.2072 | -90.4899
10.2079 | -90.4896
10.2087 | -90.4888
10.2095 | -90.4879
10.2103 | -90.4886
10.211 | -90.4883
10.2118 | -90.4867
10.2126 | -90.486
10.2134 | -90.4869
10.2141 | -90.4862
10.2149 | -90.4847
10.2157 | -90.4836
10.2165 | -90.4829
10.2172 | -90.4825
10.218 | -90.4811
10.2188 | -90.4794
10.2196 | -90.4785
10.2203 | -90.4776
10.2211 | -90.4773
10.2219 | -90.4769
10.2227 | -90.4759
10.2234 | -90.4749
10.2242 | -90.4732
10.225 | -90.4711
10.2258 | -90.469
10.2265 | -90.4674
10.2273 | -90.4663
10.2281 | -90.4658
10.2289 | -90.465
10.2296 | -90.4646
10.2304 | -90.4645
10.2312 | -90.4646
10.232 | -90.4643
10.2327 | -90.4632
10.2335 | -90.4621
10.2343 | -90.4613
10.2351 | -90.4606
10.2358 | -90.4603
10.2366 | -90.4597
10.2374 | -90.4579
10.2382 | -90.4567
10.2389 | -90.4559
10.2397 | -90.4552
10.2405 | -90.4544
10.2413 | -90.4542
10.242 | -90.4539
10.2428 | -90.4533
10.2436 | -90.4521
10.2444 | -90.4513
10.2451 | -90.452
10.2459 | -90.4523
10.2467 | -90.4507
10.2475 | -90.4491
10.2482 | -90.4486
10.249 | -90.4479
10.2498 | -90.4475
10.2506 | -90.4471
10.2513 | -90.4464
10.2521 | -90.4459
10.2529 | -90.4461
10.2536 | -90.4466
10.2544 | -90.4463
10.2552 | -90.447
10.256 | -90.4476
10.2567 | -90.4467
10.2575 | -90.446
10.2583 | -90.4451
10.2591 | -90.4435
10.2598 | -90.4429
10.2606 | -90.4422
10.2614 | -90.4412
10.2622 | -90.4405
10.2629 | -90.4402
10.2637 | -90.4395
10.2645 | -90.4383
10.2653 | -90.4372
10.266 | -90.436
10.2668 | -90.4341
10.2676 | -90.4337
10.2684 | -90.4331
10.2691 | -90.4318
10.2699 | -90.4308
10.2707 | -90.4302
10.2715 | -90.4298
10.2722 | -90.4293
10.273 | -90.4284
10.2738 | -90.4279
10.2746 | -90.428
10.2753 | -90.4278
10.2761 | -90.4262
10.2769 | -90.4242
10.2777 | -90.4233
10.2784 | -90.4223
10.2792 | -90.4218
10.28 | -90.4212
10.2808 | -90.4208
10.2815 | -90.4229
10.2823 | -90.4229
10.2831 | -90.4243
10.2839 | -90.4238
10.2846 | -90.4232
10.2854 | -90.4235
10.2862 | -90.4233
10.287 | -90.4227
10.2877 | -90.4217
10.2885 | -90.4202
10.2893 | -90.4197
10.2901 | -90.419
10.2908 | -90.4184
10.2916 | -90.4183
10.2924 | -90.4186
10.2932 | -90.4178
10.2939 | -90.417
10.2947 | -90.4159
10.2955 | -90.4152
10.2963 | -90.4147
10.297 | -90.414
10.2978 | -90.414
10.2986 | -90.4149
10.2994 | -90.4146
10.3001 | -90.4135
10.3009 | -90.4128
10.3017 | -90.4131
10.3025 | -90.4127
10.3032 | -90.4127
10.304 | -90.4129
10.3048 | -90.4131
10.3056 | -90.4124
10.3063 | -90.4115
10.3071 | -90.4106
10.3079 | -90.4113
10.3087 | -90.4107
10.3094 | -90.4106
10.3102 | -90.4101
10.311 | -90.4096
10.3118 | -90.4086
10.3125 | -90.4078
10.3133 | -90.4077
10.3141 | -90.4076
10.3149 | -90.407
10.3156 | -90.4072
10.3164 | -90.4066
10.3172 | -90.4054
10.318 | -90.4041
10.3187 | -90.4042
10.3195 | -90.4035
10.3203 | -90.4026
10.3211 | -90.4014
10.3218 | -90.4005
10.3226 | -90.4003
10.3234 | -90.4
10.3241 | -90.3982
10.3249 | -90.3967
10.3257 | -90.3961
10.3265 | -90.396
10.3272 | -90.3957
10.328 | -90.3952
10.3288 | -90.3946
10.3296 | -90.3941
10.3303 | -90.3935
10.3311 | -90.3931
10.3319 | -90.3923
10.3327 | -90.3914
10.3334 | -90.3914
10.3342 | -90.3931
10.335 | -90.394
10.3358 | -90.3941
10.3365 | -90.394
10.3373 | -90.3944
10.3381 | -90.3943
10.3389 | -90.3938
10.3396 | -90.3935
10.3404 | -90.3928
10.3412 | -90.392
10.342 | -90.3914
10.3427 | -90.3911
10.3435 | -90.3913
10.3443 | -90.3916
10.3451 | -90.3915
10.3458 | -90.3918
10.3466 | -90.3909
10.3474 | -90.3902
10.3482 | -90.3899
10.3489 | -90.3908
10.3497 | -90.3908
10.3505 | -90.391
10.3513 | -90.3905
10.352 | -90.3904
10.3528 | -90.3913
10.3536 | -90.3914
10.3544 | -90.3909
10.3551 | -90.3901
10.3559 | -90.389
10.3567 | -90.3885
10.3575 | -90.3881
10.3582 | -90.3879
10.359 | -90.3881
10.3598 | -90.3881
10.3606 | -90.3885
10.3613 | -90.3885
10.3621 | -90.3888
10.3629 | -90.3888
10.3637 | -90.3882
10.3644 | -90.3886
10.3652 | -90.3884
10.366 | -90.3875
10.3668 | -90.3873
10.3675 | -90.3871
10.3683 | -90.3868
10.3691 | -90.3867
10.3699 | -90.3861
10.3706 | -90.3869
10.3714 | -90.3883
10.3722 | -90.3889
10.373 | -90.3882
10.3737 | -90.3879
10.3745 | -90.3888
10.3753 | -90.3894
10.3761 | -90.3893
10.3768 | -90.3883
10.3776 | -90.388
10.3784 | -90.388
10.3792 | -90.3885
10.3799 | -90.3895
10.3807 | -90.3897
10.3815 | -90.3893
10.3823 | -90.3889
10.383 | -90.3888
10.3838 | -90.3888
10.3846 | -90.3894
10.3854 | -90.3905
10.3861 | -90.389
10.3869 | -90.3877
10.3877 | -90.3876
10.3885 | -90.3878
10.3892 | -90.3872
10.39 | -90.3876
10.3908 | -90.3888
10.3916 | -90.3898
10.3923 | -90.3903
10.3931 | -90.3898
10.3939 | -90.3849
10.3947 | -90.3853
10.3954 | -90.385
10.3962 | -90.3847
10.397 | -90.3865
10.3978 | -90.3883
10.3985 | -90.3895
10.3993 | -90.3898
10.4001 | -90.389
10.4008 | -90.3885
10.4016 | -90.3895
10.4024 | -90.3896
10.4032 | -90.3898
10.4039 | -90.3897
10.4047 | -90.39
10.4055 | -90.3902
10.4063 | -90.3902
10.407 | -90.3905
10.4078 | -90.3906
10.4086 | -90.3907
10.4094 | -90.3909
10.4101 | -90.3918
10.4109 | -90.3921
10.4117 | -90.3917
10.4125 | -90.3903
10.4132 | -90.3903
10.414 | -90.3903
10.4148 | -90.3901
10.4156 | -90.3905
10.4163 | -90.3924
10.4171 | -90.3927
10.4179 | -90.3931
10.4187 | -90.3927
10.4194 | -90.3929
10.4202 | -90.3933
10.421 | -90.3933
10.4218 | -90.3938
10.4225 | -90.3943
10.4233 | -90.3942
10.4241 | -90.3945
10.4249 | -90.3948
10.4256 | -90.3954
10.4264 | -90.3958
10.4272 | -90.3959
10.428 | -90.3959
10.4287 | -90.3962
10.4295 | -90.3968
10.4303 | -90.3976
10.4311 | -90.3983
10.4318 | -90.399
10.4326 | -90.3998
10.4334 | -90.4012
10.4342 | -90.4028
10.4349 | -90.4027
10.4357 | -90.401
10.4365 | -90.4007
10.4373 | -90.4021
10.438 | -90.4036
10.4388 | -90.4057
10.4396 | -90.4062
10.4404 | -90.4069
10.4411 | -90.4077
10.4419 | -90.4089
10.4427 | -90.4108
10.4435 | -90.4117
10.4442 | -90.4126
10.445 | -90.4142
10.4458 | -90.4149
10.4466 | -90.4161
10.4473 | -90.4175
10.4481 | -90.419
10.4489 | -90.42
10.4497 | -90.4207
10.4504 | -90.4216
10.4512 | -90.4224
10.452 | -90.4231
10.4528 | -90.4223
10.4535 | -90.422
10.4543 | -90.4227
10.4551 | -90.4241
10.4559 | -90.4253
10.4566 | -90.4262
10.4574 | -90.4269
10.4582 | -90.427
10.459 | -90.4275
10.4597 | -90.4285
10.4605 | -90.43
10.4613 | -90.4305
10.4621 | -90.4311
10.4628 | -90.4321
10.4636 | -90.4331
10.4644 | -90.4341
10.4652 | -90.4353
10.4659 | -90.4365
10.4667 | -90.4373
10.4675 | -90.4383
10.4683 | -90.4393
10.469 | -90.4403
10.4698 | -90.4424
10.4706 | -90.4438
10.4713 | -90.4449
10.4721 | -90.446
10.4729 | -90.4469
10.4737 | -90.4481
10.4744 | -90.4498
10.4752 | -90.4512
10.476 | -90.452
10.4768 | -90.4527
10.4775 | -90.454
10.4783 | -90.4548
10.4791 | -90.4561
10.4799 | -90.4576
10.4806 | -90.4583
10.4814 | -90.4584
10.4822 | -90.4596
10.483 | -90.4622
10.4837 | -90.4644
10.4845 | -90.4662
10.4853 | -90.4675
10.4861 | -90.4686
10.4868 | -90.4694
10.4876 | -90.4705
10.4884 | -90.4717
10.4892 | -90.4732
10.4899 | -90.4744
10.4907 | -90.4761
10.4915 | -90.4777
10.4923 | -90.4789
10.493 | -90.4794
10.4938 | -90.4804
10.4946 | -90.4827
10.4954 | -90.4841
10.4961 | -90.4859
10.4969 | -90.4887
10.4977 | -90.4899
10.4985 | -90.4914
10.4992 | -90.4931
10.5 | -90.494
10.5008 | -90.4951
10.5016 | -90.497
10.5023 | -90.4977
10.5031 | -90.4991
10.5039 | -90.5012
10.5047 | -90.5027
10.5054 | -90.5058
10.5062 | -90.5073
10.507 | -90.5088
10.5078 | -90.5114
10.5085 | -90.5146
10.5093 | -90.5163
10.5101 | -90.5192
10.5109 | -90.5232
10.5116 | -90.5256
10.5124 | -90.5266
10.5132 | -90.5291
10.514 | -90.5313
10.5147 | -90.5328
10.5155 | -90.5354
10.5163 | -90.5366
10.5171 | -90.539
10.5178 | -90.5404
10.5186 | -90.5421
10.5194 | -90.544
10.5202 | -90.5459
10.5209 | -90.5477
10.5217 | -90.5506
10.5225 | -90.553
10.5233 | -90.5562
10.524 | -90.5589
10.5248 | -90.5597
10.5256 | -90.5618
10.5264 | -90.5639
10.5271 | -90.5664
10.5279 | -90.5693
10.5287 | -90.5717
10.5295 | -90.5734
10.5302 | -90.5753
10.531 | -90.5772
10.5318 | -90.579
10.5326 | -90.5811
10.5333 | -90.5834
10.5341 | -90.5855
10.5349 | -90.5872
10.5357 | -90.5892
10.5364 | -90.5911
10.5372 | -90.593
10.538 | -90.595
10.5388 | -90.5975
10.5395 | -90.6003
10.5403 | -90.6037
10.5411 | -90.6064
10.5419 | -90.6079
10.5426 | -90.6096
10.5434 | -90.6116
10.5442 | -90.6146
10.545 | -90.6165
10.5457 | -90.6191
10.5465 | -90.6208
10.5473 | -90.6229
10.548 | -90.6251
10.5488 | -90.6278
10.5496 | -90.6302
10.5504 | -90.6336
10.5511 | -90.6363
10.5519 | -90.6383
10.5527 | -90.6406
10.5535 | -90.6422
10.5542 | -90.6434
10.555 | -90.645
10.5558 | -90.6467
10.5566 | -90.6489
10.5573 | -90.651
10.5581 | -90.6532
10.5589 | -90.6562
10.5597 | -90.6585
10.5604 | -90.659
10.5612 | -90.6595
10.562 | -90.66
10.5628 | -90.6617
10.5635 | -90.6651
10.5643 | -90.668
10.5651 | -90.6703
10.5659 | -90.6726
10.5666 | -90.6754
10.5674 | -90.6785
10.5682 | -90.6807
10.569 | -90.6824
10.5697 | -90.6852
10.5705 | -90.6878
10.5713 | -90.6903
10.5721 | -90.6925
10.5728 | -90.6952
10.5736 | -90.6969
10.5744 | -90.699
10.5752 | -90.7014
10.5759 | -90.7046
10.5767 | -90.7076
10.5775 | -90.7103
10.5783 | -90.7133
10.579 | -90.7159
10.5798 | -90.7184
10.5806 | -90.722
10.5814 | -90.726
10.5821 | -90.7286
10.5829 | -90.7311
10.5837 | -90.7328
10.5845 | -90.7355
10.5852 | -90.7372
10.586 | -90.7376
10.5868 | -90.7401
10.5876 | -90.7428
10.5883 | -90.7458
10.5891 | -90.7498
10.5899 | -90.7527
10.5907 | -90.7545
10.5914 | -90.7572
10.5922 | -90.7595
10.593 | -90.7616
10.5938 | -90.7638
10.5945 | -90.7669
10.5953 | -90.7702
10.5961 | -90.773
10.5969 | -90.7754
10.5976 | -90.7776
10.5984 | -90.7795
10.5992 | -90.7818
10.6 | -90.7846
10.6007 | -90.7869
10.6015 | -90.7894
10.6023 | -90.7919
10.6031 | -90.7937
10.6038 | -90.7961
10.6046 | -90.7975
10.6054 | -90.7989
10.6062 | -90.8005
10.6069 | -90.8035
10.6077 | -90.8075
10.6085 | -90.8116
10.6093 | -90.8146
10.61 | -90.8173
10.6108 | -90.82
10.6116 | -90.8224
10.6124 | -90.825
10.6131 | -90.828
10.6139 | -90.8306
10.6147 | -90.8331
10.6155 | -90.8358
10.6162 | -90.8382
10.617 | -90.8419
10.6178 | -90.8444
10.6186 | -90.8473
10.6193 | -90.85
10.6201 | -90.8527
10.6209 | -90.8549
10.6216 | -90.8573
10.6224 | -90.8602
10.6232 | -90.8629
10.624 | -90.8654
10.6247 | -90.8677
10.6255 | -90.8707
10.6263 | -90.8739
10.6271 | -90.8766
10.6278 | -90.8791
10.6286 | -90.8815
10.6294 | -90.8843
10.6302 | -90.8873
10.6309 | -90.8906
10.6317 | -90.8932
10.6325 | -90.8965
10.6333 | -90.9005
10.634 | -90.9043
10.6348 | -90.9072
10.6356 | -90.9098
10.6364 | -90.9126
10.6371 | -90.9151
10.6379 | -90.9179
10.6387 | -90.9205
10.6395 | -90.9233
10.6402 | -90.926
10.641 | -90.9291
10.6418 | -90.9327
10.6426 | -90.9377
10.6433 | -90.9418
10.6441 | -90.9446
10.6449 | -90.9483
10.6457 | -90.95
10.6464 | -90.9523
10.6472 | -90.954
10.648 | -90.9543
10.6488 | -90.9563
10.6495 | -90.9599
10.6503 | -90.9625
10.6511 | -90.963
10.6519 | -90.9653
10.6526 | -90.9677
10.6534 | -90.9695
10.6542 | -90.9712
10.655 | -90.9729
10.6557 | -90.9754
10.6565 | -90.9783
10.6573 | -90.9802
10.6581 | -90.981
10.6588 | -90.9831
10.6596 | -90.9858
10.6604 | -90.991
10.6612 | -90.9959
10.6619 | -90.9984
10.6627 | -91.0018
10.6635 | -91.005
10.6643 | -91.0077
10.665 | -91.0106
10.6658 | -91.0131
10.6666 | -91.0156
10.6674 | -91.0178
10.6681 | -91.0211
10.6689 | -91.0233
10.6697 | -91.0254
10.6705 | -91.0278
10.6712 | -91.0298
10.672 | -91.0312
10.6728 | -91.0319
10.6736 | -91.033
10.6743 | -91.0354
10.6751 | -91.0388
10.6759 | -91.0418
10.6767 | -91.0446
10.6774 | -91.0477
10.6782 | -91.0511
10.679 | -91.0543
10.6798 | -91.0556
10.6805 | -91.0581
10.6813 | -91.0606
10.6821 | -91.0636
10.6829 | -91.0663
10.6836 | -91.0684
10.6844 | -91.0699
10.6852 | -91.0726
10.686 | -91.0754
10.6867 | -91.078
10.6875 | -91.081
10.6883 | -91.0846
10.6891 | -91.0877
10.6898 | -91.09
10.6906 | -91.0933
10.6914 | -91.0955
10.6921 | -91.098
10.6929 | -91.1004
10.6937 | -91.1031
10.6945 | -91.1058
10.6952 | -91.1101
10.696 | -91.1145
10.6968 | -91.1178
10.6976 | -91.122
10.6983 | -91.127
10.6991 | -91.1308
10.6999 | -91.1332
10.7007 | -91.1356
10.7014 | -91.1379
10.7022 | -91.1399
10.703 | -91.1422
10.7038 | -91.1448
10.7045 | -91.148
10.7053 | -91.1507
10.7061 | -91.153
10.7069 | -91.155
10.7076 | -91.1551
10.7084 | -91.1558
10.7092 | -91.1585
10.71 | -91.1609
10.7107 | -91.1624
10.7115 | -91.1663
10.7123 | -91.173
10.7131 | -91.1785
10.7138 | -91.1816
10.7146 | -91.1842
10.7154 | -91.1865
10.7162 | -91.1895
10.7169 | -91.1925
10.7177 | -91.1942
10.7185 | -91.1957
10.7193 | -91.1975
10.72 | -91.1997
10.7208 | -91.2018
10.7216 | -91.2038
10.7224 | -91.2055
10.7231 | -91.2081
10.7239 | -91.2123
10.7247 | -91.2149
10.7255 | -91.2173
10.7262 | -91.2197
10.727 | -91.2226
10.7278 | -91.2256
10.7286 | -91.2274
10.7293 | -91.2301
10.7301 | -91.2326
10.7309 | -91.2351
10.7317 | -91.2379
10.7324 | -91.2403
10.7332 | -91.242
10.734 | -91.2425
10.7348 | -91.2442
10.7355 | -91.246
10.7363 | -91.2478
10.7371 | -91.2499
10.7379 | -91.2522
10.7386 | -91.2547
10.7394 | -91.2568
10.7402 | -91.259
10.741 | -91.2612
10.7417 | -91.2638
10.7425 | -91.2667
10.7433 | -91.2695
10.7441 | -91.2719
10.7448 | -91.2739
10.7456 | -91.2756
10.7464 | -91.2774
10.7472 | -91.2797
10.7479 | -91.2816
10.7487 | -91.2838
10.7495 | -91.286
10.7503 | -91.2886
10.751 | -91.2918
10.7518 | -91.2942
10.7526 | -91.2973
10.7534 | -91.2997
10.7541 | -91.3017
10.7549 | -91.3042
10.7557 | -91.306
10.7565 | -91.3079
10.7572 | -91.3097
10.758 | -91.3123
10.7588 | -91.3157
10.7596 | -91.3183
10.7603 | -91.3219
10.7611 | -91.325
10.7619 | -91.3268
10.7627 | -91.3288
10.7634 | -91.3316
10.7642 | -91.3336
10.765 | -91.3359
10.7658 | -91.3382
10.7665 | -91.3404
10.7673 | -91.343
10.7681 | -91.3454
10.7688 | -91.3477
10.7696 | -91.3516
10.7704 | -91.3573
10.7712 | -91.3616
10.7719 | -91.3638
10.7727 | -91.3661
10.7735 | -91.3682
10.7743 | -91.3706
10.775 | -91.373
10.7758 | -91.3749
10.7766 | -91.3775
10.7774 | -91.3804
10.7781 | -91.3818
10.7789 | -91.3834
10.7797 | -91.3851
10.7805 | -91.387
10.7812 | -91.3888
10.782 | -91.3911
10.7828 | -91.3927
10.7836 | -91.3946
10.7843 | -91.3968
10.7851 | -91.3995
10.7859 | -91.4011
10.7867 | -91.4029
10.7874 | -91.4048
10.7882 | -91.4065
10.789 | -91.4088
10.7898 | -91.4112
10.7905 | -91.4132
10.7913 | -91.4149
10.7921 | -91.4162
10.7929 | -91.4184
10.7936 | -91.4203
10.7944 | -91.4232
10.7952 | -91.4272
10.796 | -91.4288
10.7967 | -91.4326
10.7975 | -91.435
10.7983 | -91.4371
10.7991 | -91.4387
10.7998 | -91.4402
10.8006 | -91.4419
10.8014 | -91.443
10.8022 | -91.4458
10.8029 | -91.4478
10.8037 | -91.4495
10.8045 | -91.4511
10.8053 | -91.4529
10.806 | -91.4547
10.8068 | -91.4567
10.8076 | -91.4595
10.8084 | -91.4621
10.8091 | -91.4638
10.8099 | -91.4645
10.8107 | -91.4645
10.8115 | -91.4659
10.8122 | -91.469
10.813 | -91.4721
10.8138 | -91.473
10.8146 | -91.4741
10.8153 | -91.475
10.8161 | -91.4765
10.8169 | -91.4788
10.8177 | -91.4824
10.8184 | -91.4823
10.8192 | -91.4852
10.82 | -91.487
10.8208 | -91.4882
10.8215 | -91.4891
10.8223 | -91.4906
10.8231 | -91.4931
10.8239 | -91.495
10.8246 | -91.498
10.8254 | -91.5019
10.8262 | -91.5051
10.827 | -91.5073
10.8277 | -91.51
10.8285 | -91.5125
10.8293 | -91.5144
10.8301 | -91.5161
10.8308 | -91.5175
10.8316 | -91.5194
10.8324 | -91.5209
10.8332 | -91.5225
10.8339 | -91.5242
10.8347 | -91.5261
10.8355 | -91.5283
10.8363 | -91.5316
10.837 | -91.5332
10.8378 | -91.5331
10.8386 | -91.5345
10.8393 | -91.5364
10.8401 | -91.5385
10.8409 | -91.5401
10.8417 | -91.5413
10.8424 | -91.5432
10.8432 | -91.5458
10.844 | -91.5467
10.8448 | -91.5483
10.8455 | -91.5497
10.8463 | -91.5507
10.8471 | -91.554
10.8479 | -91.556
10.8486 | -91.557
10.8494 | -91.5576
10.8502 | -91.56
10.851 | -91.5625
10.8517 | -91.564
10.8525 | -91.5662
10.8533 | -91.5702
10.8541 | -91.5723
10.8548 | -91.5735
10.8556 | -91.5757
10.8564 | -91.578
10.8572 | -91.5805
10.8579 | -91.5821
10.8587 | -91.5833
10.8595 | -91.5851
10.8603 | -91.5876
10.861 | -91.5908
10.8618 | -91.5934
10.8626 | -91.5946
10.8634 | -91.5955
10.8641 | -91.5967
10.8649 | -91.5981
10.8657 | -91.5998
10.8665 | -91.5997
10.8672 | -91.6004
10.868 | -91.6006
10.8688 | -91.6011
10.8696 | -91.6025
10.8703 | -91.6045
10.8711 | -91.6053
10.8719 | -91.6061
10.8727 | -91.6072
10.8734 | -91.6074
10.8742 | -91.6077
10.875 | -91.6075
10.8758 | -91.6083
10.8765 | -91.6107
10.8773 | -91.6126
10.8781 | -91.6133
10.8789 | -91.6151
10.8796 | -91.6168
10.8804 | -91.6184
10.8812 | -91.62
10.882 | -91.6214
10.8827 | -91.6247
10.8835 | -91.6278
10.8843 | -91.6297
10.8851 | -91.6313
10.8858 | -91.6328
10.8866 | -91.6337
10.8874 | -91.6364
10.8882 | -91.6393
10.8889 | -91.6418
10.8897 | -91.6438
10.8905 | -91.6453
10.8913 | -91.6473
10.892 | -91.6484
10.8928 | -91.6498
10.8936 | -91.651
10.8944 | -91.6524
10.8951 | -91.6545
10.8959 | -91.6554
10.8967 | -91.6567
10.8975 | -91.658
10.8982 | -91.659
10.899 | -91.6603
10.8998 | -91.6624
10.9006 | -91.6633
10.9013 | -91.6645
10.9021 | -91.6655
10.9029 | -91.6661
10.9037 | -91.6673
10.9044 | -91.6685
10.9052 | -91.6695
10.906 | -91.6693
10.9068 | -91.6701
10.9075 | -91.6717
10.9083 | -91.673
10.9091 | -91.6739
10.9099 | -91.6749
10.9106 | -91.6757
10.9114 | -91.6767
10.9122 | -91.6777
10.913 | -91.6782
10.9137 | -91.6806
10.9145 | -91.6818
10.9153 | -91.6825
10.916 | -91.6834
10.9168 | -91.6842
10.9176 | -91.6857
10.9184 | -91.6874
10.9191 | -91.6893
10.9199 | -91.6905
10.9207 | -91.6924
10.9215 | -91.6954
10.9222 | -91.6971
10.923 | -91.6988
10.9238 | -91.7005
10.9246 | -91.7021
10.9253 | -91.7041
10.9261 | -91.7044
10.9269 | -91.7046
10.9277 | -91.7038
10.9284 | -91.7041
10.9292 | -91.7046
10.93 | -91.7037
10.9308 | -91.7037
10.9315 | -91.7046
10.9323 | -91.7042
10.9331 | -91.7033
10.9339 | -91.7045
10.9346 | -91.706
10.9354 | -91.7069
10.9362 | -91.7078
10.937 | -91.7089
10.9377 | -91.7094
10.9385 | -91.71
10.9393 | -91.7109
10.9401 | -91.7122
10.9408 | -91.7127
10.9416 | -91.7122
10.9424 | -91.7134
10.9432 | -91.7147
10.9439 | -91.7163
10.9447 | -91.7178
10.9455 | -91.7192
10.9463 | -91.7212
10.947 | -91.723
10.9478 | -91.7239
10.9486 | -91.7258
10.9494 | -91.7283
10.9501 | -91.7297
10.9509 | -91.7319
10.9517 | -91.7342
10.9525 | -91.7357
10.9532 | -91.737
10.954 | -91.7384
10.9548 | -91.7399
10.9556 | -91.7403
10.9563 | -91.7413
10.9571 | -91.7433
10.9579 | -91.7454
10.9587 | -91.7463
10.9594 | -91.7476
10.9602 | -91.7502
10.961 | -91.7528
10.9618 | -91.7553
10.9625 | -91.7577
10.9633 | -91.7602
10.9641 | -91.761
10.9649 | -91.7623
10.9656 | -91.7631
10.9664 | -91.764
10.9672 | -91.7651
10.968 | -91.7658
10.9687 | -91.7664
10.9695 | -91.7664
10.9703 | -91.7663
10.9711 | -91.7666
10.9718 | -91.7668
10.9726 | -91.767
10.9734 | -91.767
10.9742 | -91.7686
10.9749 | -91.77
10.9757 | -91.7718
10.9765 | -91.7727
10.9773 | -91.7727
10.978 | -91.7734
10.9788 | -91.7736
10.9796 | -91.7739
10.9804 | -91.774
10.9811 | -91.7742
10.9819 | -91.7731
10.9827 | -91.772
10.9835 | -91.7732
10.9842 | -91.7754
10.985 | -91.7779
10.9858 | -91.7789
10.9866 | -91.7803
10.9873 | -91.7808
10.9881 | -91.7813
10.9889 | -91.782
10.9896 | -91.7825
10.9904 | -91.7829
10.9912 | -91.7839
10.992 | -91.7846
10.9927 | -91.7852
10.9935 | -91.7862
10.9943 | -91.7875
10.9951 | -91.7881
10.9958 | -91.7889
10.9966 | -91.7897
10.9974 | -91.7902
10.9982 | -91.7912
10.9989 | -91.791
10.9997 | -91.79
11.0005 | -91.79
11.0013 | -91.7905
11.002 | -91.7905
11.0028 | -91.7907
11.0036 | -91.7919
11.0044 | -91.793
11.0051 | -91.7942
11.0059 | -91.7967
11.0067 | -91.7983
11.0075 | -91.7996
11.0082 | -91.8007
11.009 | -91.8018
11.0098 | -91.8024
11.0106 | -91.8029
11.0113 | -91.8035
11.0121 | -91.8044
11.0129 | -91.8047
11.0137 | -91.804
11.0144 | -91.8041
11.0152 | -91.8059
11.016 | -91.8083
11.0168 | -91.8093
11.0175 | -91.8101
11.0183 | -91.8109
11.0191 | -91.8114
11.0199 | -91.8127
11.0206 | -91.8137
11.0214 | -91.8143
11.0222 | -91.8152
11.023 | -91.8159
11.0237 | -91.8165
11.0245 | -91.8173
11.0253 | -91.8181
11.0261 | -91.8192
11.0268 | -91.8201
11.0276 | -91.821
11.0284 | -91.8228
11.0292 | -91.8246
11.0299 | -91.8254
11.0307 | -91.8268
11.0315 | -91.8285
11.0323 | -91.8302
11.033 | -91.8312
11.0338 | -91.8324
11.0346 | -91.8329
11.0354 | -91.8329
11.0361 | -91.8326
11.0369 | -91.8332
11.0377 | -91.8336
11.0385 | -91.8338
11.0392 | -91.8347
11.04 | -91.8354
11.0408 | -91.8357
11.0416 | -91.8357
11.0423 | -91.8363
11.0431 | -91.8376
11.0439 | -91.839
11.0447 | -91.8398
11.0454 | -91.8406
11.0462 | -91.8412
11.047 | -91.8419
11.0478 | -91.8421
11.0485 | -91.8426
11.0493 | -91.844
11.0501 | -91.8446
11.0509 | -91.8458
11.0516 | -91.8465
11.0524 | -91.8468
11.0532 | -91.8474
11.054 | -91.8488
11.0547 | -91.8509
11.0555 | -91.852
11.0563 | -91.8528
11.0571 | -91.8541
11.0578 | -91.8548
11.0586 | -91.8549
11.0594 | -91.8564
11.0602 | -91.8578
11.0609 | -91.8584
11.0617 | -91.8593
11.0625 | -91.8606
11.0633 | -91.8623
11.064 | -91.8628
11.0648 | -91.8631
11.0656 | -91.8627
11.0663 | -91.8645
11.0671 | -91.8677
11.0679 | -91.8699
11.0687 | -91.8712
11.0694 | -91.8723
11.0702 | -91.8731
11.071 | -91.8738
11.0718 | -91.8741
11.0725 | -91.875
11.0733 | -91.8759
11.0741 | -91.8764
11.0749 | -91.877
11.0756 | -91.8783
11.0764 | -91.8798
11.0772 | -91.8808
11.078 | -91.8813
11.0787 | -91.8814
11.0795 | -91.8819
11.0803 | -91.8825
11.0811 | -91.8834
11.0818 | -91.884
11.0826 | -91.8852
11.0834 | -91.8863
11.0842 | -91.8875
11.0849 | -91.8889
11.0857 | -91.8899
11.0865 | -91.8914
11.0873 | -91.8924
11.088 | -91.8929
11.0888 | -91.8936
11.0896 | -91.8939
11.0904 | -91.8955
11.0911 | -91.897
11.0919 | -91.8979
11.0927 | -91.8993
11.0935 | -91.9
11.0942 | -91.9006
11.095 | -91.9018
11.0958 | -91.9029
11.0966 | -91.9039
11.0973 | -91.9051
11.0981 | -91.9053
11.0989 | -91.9056
11.0997 | -91.907
11.1004 | -91.9083
11.1012 | -91.91
11.102 | -91.9109
11.1028 | -91.9125
11.1035 | -91.9146
11.1043 | -91.9159
11.1051 | -91.9174
11.1059 | -91.9181
11.1066 | -91.9186
11.1074 | -91.9177
11.1082 | -91.9188
11.109 | -91.9204
11.1097 | -91.9217
11.1105 | -91.9222
11.1113 | -91.9235
11.1121 | -91.9251
11.1128 | -91.9264
11.1136 | -91.9276
11.1144 | -91.9285
11.1152 | -91.9294
11.1159 | -91.9313
11.1167 | -91.9325
11.1175 | -91.9334
11.1183 | -91.9343
11.119 | -91.9357
11.1198 | -91.9368
11.1206 | -91.938
11.1214 | -91.9384
11.1221 | -91.9391
11.1229 | -91.9402
11.1237 | -91.9421
11.1245 | -91.9423
11.1252 | -91.9423
11.126 | -91.9429
11.1268 | -91.9439
11.1276 | -91.9447
11.1283 | -91.9457
11.1291 | -91.9465
11.1299 | -91.947
11.1307 | -91.9483
11.1314 | -91.9481
11.1322 | -91.9485
11.133 | -91.9508
11.1338 | -91.9545
11.1345 | -91.9568
11.1353 | -91.9572
11.1361 | -91.9578
11.1369 | -91.9594
11.1376 | -91.9607
11.1384 | -91.9625
11.1392 | -91.9651
11.1399 | -91.9661
11.1407 | -91.9675
11.1415 | -91.9712
11.1423 | -91.9769
11.143 | -91.9772
11.1438 | -91.9785
11.1446 | -91.98
11.1454 | -91.9804
11.1461 | -91.9808
11.1469 | -91.9816
11.1477 | -91.9825
11.1485 | -91.9838
11.1492 | -91.9845
11.15 | -91.9847
11.1508 | -91.985
11.1516 | -91.9845
11.1523 | -91.9846
11.1531 | -91.9848
11.1539 | -91.9853
11.1547 | -91.9855
11.1554 | -91.9848
11.1562 | -91.9848
11.157 | -91.9855
11.1578 | -91.9865
11.1585 | -91.9872
11.1593 | -91.9876
11.1601 | -91.989
11.1609 | -91.991
11.1616 | -91.9941
11.1624 | -91.9953
11.1632 | -91.9966
11.164 | -91.9982
11.1647 | -91.9988
11.1655 | -91.9993
11.1663 | -91.9998
11.1671 | -92.0008
11.1678 | -92.0017
11.1686 | -92.0034
11.1694 | -92.0045
11.1702 | -92.0051
11.1709 | -92.005
11.1717 | -92.0053
11.1725 | -92.0065
11.1733 | -92.008
11.174 | -92.0083
11.1748 | -92.0078
11.1756 | -92.0082
11.1764 | -92.0089
11.1771 | -92.0095
11.1779 | -92.0107
11.1787 | -92.0124
11.1795 | -92.0134
11.1802 | -92.014
11.181 | -92.015
11.1818 | -92.0159
11.1826 | -92.0153
11.1833 | -92.0143
11.1841 | -92.0142
11.1849 | -92.0115
11.1857 | -92.0125
11.1864 | -92.0164
11.1872 | -92.0183
11.188 | -92.0189
11.1888 | -92.02
11.1895 | -92.0209
11.1903 | -92.0217
11.1911 | -92.0228
11.1919 | -92.0246
11.1926 | -92.0261
11.1934 | -92.0271
11.1942 | -92.0275
11.195 | -92.0281
11.1957 | -92.0284
11.1965 | -92.0289
11.1973 | -92.0303
11.1981 | -92.0321
11.1988 | -92.0325
11.1996 | -92.0342
11.2004 | -92.0361
11.2012 | -92.0369
11.2019 | -92.0379
11.2027 | -92.0388
11.2035 | -92.0396
11.2043 | -92.04
11.205 | -92.0422
11.2058 | -92.0446
11.2066 | -92.0454
11.2074 | -92.046
11.2081 | -92.0469
11.2089 | -92.0475
11.2097 | -92.0488
11.2104 | -92.0509
11.2112 | -92.0524
11.212 | -92.0527
11.2128 | -92.053
11.2135 | -92.0548
11.2143 | -92.0564
11.2151 | -92.0574
11.2159 | -92.0579
11.2166 | -92.0583
11.2174 | -92.0589
11.2182 | -92.058
11.219 | -92.0577
11.2197 | -92.0583
11.2205 | -92.0579
11.2213 | -92.0588
11.2221 | -92.0599
11.2228 | -92.0613
11.2236 | -92.0621
11.2244 | -92.0628
11.2252 | -92.063
11.2259 | -92.063
11.2267 | -92.0633
11.2275 | -92.065
11.2283 | -92.0647
11.229 | -92.0633
11.2298 | -92.0637
11.2306 | -92.0646
11.2314 | -92.0645
11.2321 | -92.0654
11.2329 | -92.0662
11.2337 | -92.0678
11.2345 | -92.0692
11.2352 | -92.0721
11.236 | -92.0738
11.2368 | -92.0745
11.2376 | -92.0756
11.2383 | -92.0766
11.2391 | -92.0775
11.2399 | -92.0783
11.2407 | -92.0796
11.2414 | -92.0812
11.2422 | -92.0831
11.243 | -92.0848
11.2438 | -92.0861
11.2445 | -92.0869
11.2453 | -92.0879
11.2461 | -92.0884
11.2469 | -92.0885
11.2476 | -92.0904
11.2484 | -92.0926
11.2492 | -92.094
11.25 | -92.095
11.2507 | -92.0961
11.2515 | -92.0969
11.2523 | -92.0981
11.2531 | -92.0991
11.2538 | -92.099
11.2546 | -92.0985
11.2554 | -92.1006
11.2562 | -92.1023
11.2569 | -92.103
11.2577 | -92.1048
11.2585 | -92.1072
11.2593 | -92.1086
11.26 | -92.1096
11.2608 | -92.1098
11.2616 | -92.109
11.2624 | -92.1089
11.2631 | -92.109
11.2639 | -92.1092
11.2647 | -92.1096
11.2655 | -92.1097
11.2662 | -92.1104
11.267 | -92.1108
11.2678 | -92.1104
11.2686 | -92.1105
11.2693 | -92.1111
11.2701 | -92.112
11.2709 | -92.113
11.2717 | -92.1143
11.2724 | -92.1139
11.2732 | -92.1149
11.274 | -92.1144
11.2748 | -92.1143
11.2755 | -92.1149
11.2763 | -92.1159
11.2771 | -92.1165
11.2779 | -92.1171
11.2786 | -92.1181
11.2794 | -92.1193
11.2802 | -92.1199
11.281 | -92.1209
11.2817 | -92.123
11.2825 | -92.124
11.2833 | -92.125
11.2841 | -92.1267
11.2848 | -92.1279
11.2856 | -92.1279
11.2864 | -92.1278
11.2871 | -92.1286
11.2879 | -92.1294
11.2887 | -92.1295
11.2895 | -92.1303
11.2902 | -92.1316
11.291 | -92.132
11.2918 | -92.1326
11.2926 | -92.1336
11.2933 | -92.1346
11.2941 | -92.1355
11.2949 | -92.1364
11.2957 | -92.1373
11.2964 | -92.1386
11.2972 | -92.1395
11.298 | -92.1401
11.2988 | -92.1411
11.2995 | -92.1427
11.3003 | -92.1446
11.3011 | -92.1455
11.3019 | -92.1466
11.3026 | -92.1474
11.3034 | -92.1484
11.3042 | -92.1496
11.305 | -92.1513
11.3057 | -92.1517
11.3065 | -92.1522
11.3073 | -92.153
11.3081 | -92.1541
11.3088 | -92.1552
11.3096 | -92.1559
11.3104 | -92.1565
11.3112 | -92.1574
11.3119 | -92.1582
11.3127 | -92.1583
11.3135 | -92.1591
11.3143 | -92.1611
11.315 | -92.1627
11.3158 | -92.164
11.3166 | -92.165
11.3174 | -92.1659
11.3181 | -92.1668
11.3189 | -92.1679
11.3197 | -92.1698
11.3205 | -92.1727
11.3212 | -92.1759
11.322 | -92.179
11.3228 | -92.1806
11.3236 | -92.1812
11.3243 | -92.1816
11.3251 | -92.1809
11.3259 | -92.1819
11.3267 | -92.1827
11.3274 | -92.1837
11.3282 | -92.1845
11.329 | -92.1852
11.3298 | -92.186
11.3305 | -92.1868
11.3313 | -92.1873
11.3321 | -92.1879
11.3329 | -92.1893
11.3336 | -92.1901
11.3344 | -92.1901
11.3352 | -92.1905
11.336 | -92.1908
11.3367 | -92.1904
11.3375 | -92.1905
11.3383 | -92.192
11.3391 | -92.1934
11.3398 | -92.1943
11.3406 | -92.1946
11.3414 | -92.1945
11.3422 | -92.195
11.3429 | -92.1963
11.3437 | -92.1975
11.3445 | -92.1985
11.3453 | -92.2003
11.346 | -92.2012
11.3468 | -92.2027
11.3476 | -92.206
11.3484 | -92.2072
11.3491 | -92.2075
11.3499 | -92.2073
11.3507 | -92.2081
11.3515 | -92.2094
11.3522 | -92.2108
11.353 | -92.2113
11.3538 | -92.2126
11.3546 | -92.2125
11.3553 | -92.2129
11.3561 | -92.2149
11.3569 | -92.2169
11.3576 | -92.2174
11.3584 | -92.218
11.3592 | -92.2183
11.36 | -92.2186
11.3607 | -92.22
11.3615 | -92.2225
11.3623 | -92.224
11.3631 | -92.225
11.3638 | -92.2262
11.3646 | -92.2269
11.3654 | -92.2266
11.3662 | -92.2278
11.3669 | -92.2291
11.3677 | -92.231
11.3685 | -92.2336
11.3693 | -92.2355
11.37 | -92.2364
11.3708 | -92.2369
11.3716 | -92.2381
11.3724 | -92.2394
11.3731 | -92.2403
11.3739 | -92.2413
11.3747 | -92.2421
11.3755 | -92.2425
11.3762 | -92.2433
11.377 | -92.2445
11.3778 | -92.2453
11.3786 | -92.2465
11.3793 | -92.2471
11.3801 | -92.2468
11.3809 | -92.2464
11.3817 | -92.2464
11.3824 | -92.2463
11.3832 | -92.2466
11.384 | -92.2466
11.3848 | -92.2465
11.3855 | -92.2471
11.3863 | -92.2471
11.3871 | -92.2466
11.3879 | -92.2458
11.3886 | -92.2458
11.3894 | -92.2469
11.3902 | -92.2478
11.391 | -92.2485
11.3917 | -92.2492
11.3925 | -92.2488
11.3933 | -92.2495
11.3941 | -92.2505
11.3948 | -92.2518
11.3956 | -92.2527
11.3964 | -92.2544
11.3972 | -92.2556
11.3979 | -92.2561
11.3987 | -92.2569
11.3995 | -92.2556
11.4003 | -92.2555
11.401 | -92.256
11.4018 | -92.2565
11.4026 | -92.257
11.4034 | -92.2567
11.4041 | -92.2571
11.4049 | -92.2582
11.4057 | -92.259
11.4065 | -92.2596
11.4072 | -92.26
11.408 | -92.2604
11.4088 | -92.2612
11.4096 | -92.2622
11.4103 | -92.2637
11.4111 | -92.2645
11.4119 | -92.2655
11.4127 | -92.2661
11.4134 | -92.2675
11.4142 | -92.2712
11.415 | -92.2729
11.4158 | -92.2737
11.4165 | -92.2749
11.4173 | -92.2772
11.4181 | -92.2785
11.4189 | -92.2794
11.4196 | -92.2812
11.4204 | -92.2835
11.4212 | -92.2855
11.422 | -92.2838
11.4227 | -92.2808
11.4235 | -92.2809
11.4243 | -92.2818
11.4251 | -92.2845
11.4258 | -92.2856
11.4266 | -92.2866
11.4274 | -92.2872
11.4282 | -92.2882
11.4289 | -92.2891
11.4297 | -92.2903
11.4305 | -92.2904
11.4313 | -92.2903
11.432 | -92.2909
11.4328 | -92.2916
11.4336 | -92.2921
11.4343 | -92.2928
11.4351 | -92.2914
11.4359 | -92.29
11.4367 | -92.2917
11.4374 | -92.2925
11.4382 | -92.2932
11.439 | -92.2944
11.4398 | -92.2967
11.4405 | -92.3006
11.4413 | -92.3033
11.4421 | -92.3029
11.4429 | -92.3029
11.4436 | -92.3036
11.4444 | -92.3045
11.4452 | -92.3047
11.446 | -92.305
11.4467 | -92.305
11.4475 | -92.3052
11.4483 | -92.3059
11.4491 | -92.3062
11.4498 | -92.3066
11.4506 | -92.307
11.4514 | -92.3079
11.4522 | -92.3084
11.4529 | -92.308
11.4537 | -92.3094
11.4545 | -92.3115
11.4553 | -92.3127
11.456 | -92.313
11.4568 | -92.3132
11.4576 | -92.3123
11.4584 | -92.3124
11.4591 | -92.3128
11.4599 | -92.3138
11.4607 | -92.3152
11.4615 | -92.3158
11.4622 | -92.3162
11.463 | -92.3156
11.4638 | -92.314
11.4646 | -92.3132
11.4653 | -92.3131
11.4661 | -92.3128
11.4669 | -92.3134
11.4677 | -92.3127
11.4684 | -92.3134
11.4692 | -92.3142
11.47 | -92.3145
11.4708 | -92.3149
11.4715 | -92.315
11.4723 | -92.3146
11.4731 | -92.3146
11.4739 | -92.3147
11.4746 | -92.3151
11.4754 | -92.3161
11.4762 | -92.3171
11.477 | -92.3176
11.4777 | -92.3179
11.4785 | -92.3197
11.4793 | -92.3221
11.4801 | -92.3217
11.4808 | -92.3207
11.4816 | -92.3215
11.4824 | -92.3217
11.4832 | -92.3221
11.4839 | -92.3222
11.4847 | -92.3214
11.4855 | -92.32
11.4863 | -92.3191
11.487 | -92.3191
11.4878 | -92.3192
11.4886 | -92.3206
11.4894 | -92.3211
11.4901 | -92.322
11.4909 | -92.3225
11.4917 | -92.3226
11.4925 | -92.3227
11.4932 | -92.3225
11.494 | -92.3229
11.4948 | -92.3228
11.4956 | -92.3219
11.4963 | -92.3217
11.4971 | -92.3219
11.4979 | -92.3217
11.4987 | -92.3217
11.4994 | -92.322
11.5002 | -92.3206
11.501 | -92.3191
11.5018 | -92.3193
11.5025 | -92.3195
11.5033 | -92.3197
11.5041 | -92.3198
11.5048 | -92.3195
11.5056 | -92.3192
11.5064 | -92.3211
11.5072 | -92.3222
11.5079 | -92.3217
11.5087 | -92.3221
11.5095 | -92.3214
11.5103 | -92.3199
11.511 | -92.3194
11.5118 | -92.3192
11.5126 | -92.3195
11.5134 | -92.3191
11.5141 | -92.3187
11.5149 | -92.3182
11.5157 | -92.3182
11.5165 | -92.3182
11.5172 | -92.3183
11.518 | -92.3186
11.5188 | -92.3189
11.5196 | -92.3189
11.5203 | -92.3188
11.5211 | -92.3191
11.5219 | -92.3188
11.5227 | -92.3184
11.5234 | -92.3181
11.5242 | -92.3178
11.525 | -92.3178
11.5258 | -92.3173
11.5265 | -92.3167
11.5273 | -92.3169
11.5281 | -92.3178
11.5289 | -92.3171
11.5296 | -92.3163
11.5304 | -92.3157
11.5312 | -92.3159
11.532 | -92.3161
11.5327 | -92.3152
11.5335 | -92.3127
11.5343 | -92.3105
11.5351 | -92.3093
11.5358 | -92.309
11.5366 | -92.3095
11.5374 | -92.309
11.5382 | -92.3084
11.5389 | -92.3085
11.5397 | -92.3081
11.5405 | -92.3084
11.5413 | -92.3085
11.542 | -92.3093
11.5428 | -92.3102
11.5436 | -92.3106
11.5444 | -92.3106
11.5451 | -92.3105
11.5459 | -92.3103
11.5467 | -92.3095
11.5475 | -92.309
11.5482 | -92.309
11.549 | -92.3085
11.5498 | -92.3084
11.5506 | -92.3082
11.5513 | -92.3079
11.5521 | -92.3066
11.5529 | -92.3057
11.5537 | -92.306
11.5544 | -92.3065
11.5552 | -92.3067
11.556 | -92.3064
11.5568 | -92.3061
11.5575 | -92.3048
11.5583 | -92.3044
11.5591 | -92.3045
11.5599 | -92.3041
11.5606 | -92.3027
11.5614 | -92.3014
11.5622 | -92.3011
11.563 | -92.3
11.5637 | -92.2992
11.5645 | -92.2976
11.5653 | -92.2946
11.5661 | -92.2927
11.5668 | -92.2909
11.5676 | -92.2908
11.5684 | -92.291
11.5692 | -92.2904
11.5699 | -92.2899
11.5707 | -92.2888
11.5715 | -92.2866
11.5723 | -92.2861
11.573 | -92.2858
11.5738 | -92.285
11.5746 | -92.2837
11.5754 | -92.2828
11.5761 | -92.2823
11.5769 | -92.282
11.5777 | -92.2812
11.5785 | -92.2804
11.5792 | -92.2815
11.58 | -92.2814
11.5808 | -92.2806
11.5815 | -92.28
11.5823 | -92.2788
11.5831 | -92.2777
11.5839 | -92.2771
11.5846 | -92.2768
11.5854 | -92.2762
11.5862 | -92.2758
11.587 | -92.2756
11.5877 | -92.2764
11.5885 | -92.2773
11.5893 | -92.275
11.5901 | -92.274
11.5908 | -92.274
11.5916 | -92.2744
11.5924 | -92.274
11.5932 | -92.2726
11.5939 | -92.2712
11.5947 | -92.2705
11.5955 | -92.2689
11.5963 | -92.2682
11.597 | -92.2687
11.5978 | -92.2686
11.5986 | -92.2679
11.5994 | -92.2674
11.6001 | -92.2677
11.6009 | -92.266
11.6017 | -92.2642
11.6025 | -92.262
11.6032 | -92.2608
11.604 | -92.26
11.6048 | -92.2593
11.6056 | -92.2583
11.6063 | -92.2571
11.6071 | -92.2561
11.6079 | -92.2553
11.6087 | -92.2545
11.6094 | -92.2537
11.6102 | -92.2532
11.611 | -92.2521
11.6118 | -92.2498
11.6125 | -92.2475
11.6133 | -92.2462
11.6141 | -92.2452
11.6149 | -92.2434
11.6156 | -92.2426
11.6164 | -92.2414
11.6172 | -92.242
11.618 | -92.2415
11.6187 | -92.24
11.6195 | -92.2388
11.6203 | -92.2372
11.6211 | -92.2358
11.6218 | -92.235
11.6226 | -92.2321
11.6234 | -92.2301
11.6242 | -92.2285
11.6249 | -92.2273
11.6257 | -92.2249
11.6265 | -92.2236
11.6273 | -92.2234
11.628 | -92.2226
11.6288 | -92.2217
11.6296 | -92.2196
11.6304 | -92.2156
11.6311 | -92.2119
11.6319 | -92.2078
11.6327 | -92.2044
11.6335 | -92.2034
11.6342 | -92.2022
11.635 | -92.2012
11.6358 | -92.1996
11.6366 | -92.1979
11.6373 | -92.1963
11.6381 | -92.1947
11.6389 | -92.1933
11.6397 | -92.1922
11.6404 | -92.1909
11.6412 | -92.189
11.642 | -92.1874
11.6428 | -92.1858
11.6435 | -92.1845
11.6443 | -92.1835
11.6451 | -92.1822
11.6459 | -92.1806
11.6466 | -92.178
11.6474 | -92.1742
11.6482 | -92.1718
11.649 | -92.1723
11.6497 | -92.172
11.6505 | -92.1708
11.6513 | -92.1695
11.652 | -92.168
11.6528 | -92.1663
11.6536 | -92.1651
11.6544 | -92.1641
11.6551 | -92.161
11.6559 | -92.1581
11.6567 | -92.1573
11.6575 | -92.1568
11.6582 | -92.1538
11.659 | -92.1512
11.6598 | -92.1495
11.6606 | -92.147
11.6613 | -92.1438
11.6621 | -92.1413
11.6629 | -92.1384
11.6637 | -92.1363
11.6644 | -92.1343
11.6652 | -92.1324
11.666 | -92.131
11.6668 | -92.129
11.6675 | -92.1248
11.6683 | -92.1218
11.6691 | -92.1196
11.6699 | -92.1179
11.6706 | -92.1162
11.6714 | -92.1144
11.6722 | -92.1126
11.673 | -92.1105
11.6737 | -92.109
11.6745 | -92.1065
11.6753 | -92.1032
11.6761 | -92.1
11.6768 | -92.0979
11.6776 | -92.096
11.6784 | -92.0942
11.6792 | -92.0921
11.6799 | -92.0907
11.6807 | -92.0897
11.6815 | -92.088
11.6823 | -92.0852
11.683 | -92.081
11.6838 | -92.0783
11.6846 | -92.0769
11.6854 | -92.0749
11.6861 | -92.0723
11.6869 | -92.0697
11.6877 | -92.0677
11.6885 | -92.0655
11.6892 | -92.0638
11.69 | -92.0603
11.6908 | -92.0587
11.6916 | -92.057
11.6923 | -92.055
11.6931 | -92.0519
11.6939 | -92.0481
11.6947 | -92.0445
11.6954 | -92.0427
11.6962 | -92.0404
11.697 | -92.0368
11.6978 | -92.0339
11.6985 | -92.0315
11.6993 | -92.0288
11.7001 | -92.0273
11.7009 | -92.0249
11.7016 | -92.0212
11.7024 | -92.0182
11.7032 | -92.0159
11.704 | -92.0136
11.7047 | -92.0113
11.7055 | -92.0091
11.7063 | -92.0064
11.7071 | -92.0036
11.7078 | -92.0016
11.7086 | -91.9999
11.7094 | -91.9982
11.7102 | -91.9964
11.7109 | -91.9946
11.7117 | -91.9927
11.7125 | -91.9903
11.7133 | -91.9886
11.714 | -91.9875
11.7148 | -91.9859
11.7156 | -91.9841
11.7164 | -91.9827
11.7171 | -91.9814
11.7179 | -91.98
11.7187 | -91.9772
11.7195 | -91.9741
11.7202 | -91.9726
11.721 | -91.971
11.7218 | -91.9685
11.7226 | -91.9657
11.7233 | -91.9626
11.7241 | -91.9572
11.7249 | -91.9543
11.7256 | -91.952
11.7264 | -91.9486
11.7272 | -91.9456
11.728 | -91.9441
11.7287 | -91.943
11.7295 | -91.9405
11.7303 | -91.9381
11.7311 | -91.9356
11.7318 | -91.9329
11.7326 | -91.9304
11.7334 | -91.927
11.7342 | -91.9238
11.7349 | -91.9212
11.7357 | -91.9185
11.7365 | -91.916
11.7373 | -91.9132
11.738 | -91.9106
11.7388 | -91.9081
11.7396 | -91.905
11.7404 | -91.9015
11.7411 | -91.8979
11.7419 | -91.8944
11.7427 | -91.8915
11.7435 | -91.8896
11.7442 | -91.8871
11.745 | -91.8847
11.7458 | -91.8819
11.7466 | -91.8787
11.7473 | -91.8762
11.7481 | -91.8735
11.7489 | -91.8701
11.7497 | -91.8672
11.7504 | -91.8648
11.7512 | -91.8625
11.752 | -91.8596
11.7528 | -91.857
11.7535 | -91.8544
11.7543 | -91.8516
11.7551 | -91.8487
11.7559 | -91.8462
11.7566 | -91.8438
11.7574 | -91.8409
11.7582 | -91.8393
11.759 | -91.838
11.7597 | -91.8357
11.7605 | -91.8336
11.7613 | -91.8318
11.7621 | -91.8303
11.7628 | -91.8291
11.7636 | -91.8268
11.7644 | -91.8243
11.7652 | -91.8216
11.7659 | -91.819
11.7667 | -91.8166
11.7675 | -91.8146
11.7683 | -91.8121
11.769 | -91.8089
11.7698 | -91.8055
11.7706 | -91.803
11.7714 | -91.8006
11.7721 | -91.7999
11.7729 | -91.7988
11.7737 | -91.795
11.7745 | -91.7908
11.7752 | -91.787
11.776 | -91.7837
11.7768 | -91.7806
11.7776 | -91.7781
11.7783 | -91.7755
11.7791 | -91.7728
11.7799 | -91.7695
11.7807 | -91.7655
11.7814 | -91.7624
11.7822 | -91.7586
11.783 | -91.7556
11.7838 | -91.7527
11.7845 | -91.7498
11.7853 | -91.7472
11.7861 | -91.7443
11.7869 | -91.7416
11.7876 | -91.7397
11.7884 | -91.7378
11.7892 | -91.735
11.79 | -91.7319
11.7907 | -91.7299
11.7915 | -91.7288
11.7923 | -91.7264
11.7931 | -91.7239
11.7938 | -91.7218
11.7946 | -91.7184
11.7954 | -91.7142
11.7962 | -91.711
11.7969 | -91.708
11.7977 | -91.7046
11.7985 | -91.701
11.7993 | -91.6979
11.8 | -91.6954
11.8008 | -91.6925
11.8016 | -91.69
11.8023 | -91.6879
11.8031 | -91.6852
11.8039 | -91.6828
11.8047 | -91.6804
11.8054 | -91.678
11.8062 | -91.6765
11.807 | -91.6744
11.8078 | -91.6719
11.8085 | -91.6692
11.8093 | -91.6664
11.8101 | -91.6638
11.8109 | -91.6611
11.8116 | -91.658
11.8124 | -91.6551
11.8132 | -91.6524
11.814 | -91.6493
11.8147 | -91.6472
11.8155 | -91.6433
11.8163 | -91.6393
11.8171 | -91.636
11.8178 | -91.6332
11.8186 | -91.6306
11.8194 | -91.628
11.8202 | -91.6249
11.8209 | -91.6217
11.8217 | -91.6193
11.8225 | -91.6165
11.8233 | -91.6137
11.824 | -91.6104
11.8248 | -91.6075
11.8256 | -91.6053
11.8264 | -91.6021
11.8271 | -91.5987
11.8279 | -91.5962
11.8287 | -91.5939
11.8295 | -91.5925
11.8302 | -91.591
11.831 | -91.5888
11.8318 | -91.5867
11.8326 | -91.5833
11.8333 | -91.5779
11.8341 | -91.5753
11.8349 | -91.5734
11.8357 | -91.5717
11.8364 | -91.5695
11.8372 | -91.567
11.838 | -91.5644
11.8388 | -91.5617
11.8395 | -91.5587
11.8403 | -91.5554
11.8411 | -91.5529
11.8419 | -91.5505
11.8426 | -91.5486
11.8434 | -91.5467
11.8442 | -91.5435
11.845 | -91.5411
11.8457 | -91.5386
11.8465 | -91.5367
11.8473 | -91.5347
11.8481 | -91.5326
11.8488 | -91.5297
11.8496 | -91.5275
11.8504 | -91.5248
11.8512 | -91.5223
11.8519 | -91.52
11.8527 | -91.5174
11.8535 | -91.514
11.8543 | -91.5101
11.855 | -91.5066
11.8558 | -91.5042
11.8566 | -91.5023
11.8574 | -91.5002
11.8581 | -91.498
11.8589 | -91.4962
11.8597 | -91.4932
11.8605 | -91.4905
11.8612 | -91.4894
11.862 | -91.4883
11.8628 | -91.4857
11.8636 | -91.4836
11.8643 | -91.4807
11.8651 | -91.4779
11.8659 | -91.4752
11.8667 | -91.4729
11.8674 | -91.4702
11.8682 | -91.4674
11.869 | -91.4634
11.8698 | -91.46
11.8705 | -91.4568
11.8713 | -91.4542
11.8721 | -91.4512
11.8728 | -91.4484
11.8736 | -91.4459
11.8744 | -91.4437
11.8752 | -91.4422
11.8759 | -91.4406
11.8767 | -91.4388
11.8775 | -91.4362
11.8783 | -91.4327
11.879 | -91.4298
11.8798 | -91.4272
11.8806 | -91.4251
11.8814 | -91.4225
11.8821 | -91.4193
11.8829 | -91.4168
11.8837 | -91.4143
11.8845 | -91.4117
11.8852 | -91.409
11.886 | -91.4061
11.8868 | -91.4034
11.8876 | -91.4013
11.8883 | -91.3982
11.8891 | -91.3952
11.8899 | -91.3927
11.8907 | -91.3894
11.8914 | -91.3883
11.8922 | -91.3856
11.893 | -91.3812
11.8938 | -91.3773
11.8945 | -91.3743
11.8953 | -91.3719
11.8961 | -91.3698
11.8969 | -91.3679
11.8976 | -91.3662
11.8984 | -91.3642
11.8992 | -91.3621
11.9 | -91.3598
11.9007 | -91.3573
11.9015 | -91.3549
11.9023 | -91.353
11.9031 | -91.3514
11.9038 | -91.35
11.9046 | -91.3484
11.9054 | -91.3467
11.9062 | -91.3459
11.9069 | -91.3441
11.9077 | -91.3416
11.9085 | -91.3389
11.9093 | -91.3366
11.91 | -91.3349
11.9108 | -91.3327
11.9116 | -91.3304
11.9124 | -91.328
11.9131 | -91.3256
11.9139 | -91.3249
11.9147 | -91.3229
11.9155 | -91.3205
11.9162 | -91.3178
11.917 | -91.3155
11.9178 | -91.3131
11.9186 | -91.3106
11.9193 | -91.3088
11.9201 | -91.3074
11.9209 | -91.3052
11.9217 | -91.3024
11.9224 | -91.3008
11.9232 | -91.2994
11.924 | -91.2978
11.9248 | -91.2965
11.9255 | -91.2947
11.9263 | -91.2927
11.9271 | -91.2901
11.9279 | -91.2883
11.9286 | -91.2869
11.9294 | -91.2856
11.9302 | -91.2843
11.931 | -91.2825
11.9317 | -91.2811
11.9325 | -91.2802
11.9333 | -91.2792
11.9341 | -91.2768
11.9348 | -91.2746
11.9356 | -91.2727
11.9364 | -91.2706
11.9372 | -91.2691
11.9379 | -91.2674
11.9387 | -91.2657
11.9395 | -91.264
11.9403 | -91.2623
11.941 | -91.2612
11.9418 | -91.26
11.9426 | -91.2577
11.9434 | -91.256
11.9441 | -91.2551
11.9449 | -91.254
11.9457 | -91.2522
11.9465 | -91.2503
11.9472 | -91.2489
11.948 | -91.2471
11.9488 | -91.2447
11.9495 | -91.243
11.9503 | -91.2416
11.9511 | -91.2412
11.9519 | -91.2399
11.9526 | -91.238
11.9534 | -91.2368
11.9542 | -91.2355
11.955 | -91.2343
11.9557 | -91.2335
11.9565 | -91.2321
11.9573 | -91.231
11.9581 | -91.2301
11.9588 | -91.2283
11.9596 | -91.2265
11.9604 | -91.2249
11.9612 | -91.2237
11.9619 | -91.2221
11.9627 | -91.2207
11.9635 | -91.2187
11.9643 | -91.2171
11.965 | -91.2159
11.9658 | -91.2146
11.9666 | -91.2134