## CALIBRATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 68 |  | CTBTO
Transducer Block | 02-90C/P | 65 |  | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
Hood | 02-92B/P | 54
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE REFERENCE
Atmospheric pressure | 1011.4 | hPa |  | PR 000 150_-_CAL
Humidity | 19.1 | %/Hr |  | Measurement device
Internal Temperature of Calibration Bench | 22.78 | °C |  | IS CALIBRATOR
AUX Sensor pressure | 1018.4 | hPa |  | Sensor SoH at test start
AUX Sensor Temperature | 23.59 | °C |  | OK
SENSITIVITY AT 1 Hz
Output Parameters | Theoretical Value | Measured Value @1Hz | Margin (±) | Unit | Result
Sensitivity of Pressure Derivative output | 2 | 2 | 0.04 | mV/Pa/sec | OK | Mode d'emploi pour créer un constat de vérification
Sensitivity of Pressure output | 20 | 19.97 | 0.4 | mV/Pa | OK | Cliquer sur "DEMARRER LA SAISIE"
Sensitivity of Calibration coil | 6 | 5.98 | 0.13 | Pa/V | OK | Suivre les instructions affichées
Une fois la partie SENSOR IDENTIFICATION remplie, renseigner la partie ENVIRONNEMENTAL PARAMETERS (les données environnementales)
TRANSFER FUNCTION | [5 empty] | Renseigner la partie SENSITIVITY AT 1Hz :les sensibilités en voie de dérivée de Pression (Pressure Derivative), en voie de pression (Pressure) et de Calibration (calibration Coil)
Electrical |  | Acoustical |  |  |  | Une sauvegarde temporaire est toujours possible si le fichier doit être modifié (Fichier "enregistrer")
Vérifier toutes les données saisie notamment la partie CUSTOMER ("Client") et la partie VISA (Noms des personnes et date d'ajustage)
Cliquer sur "ENREGISTREMENT DEFINITIF" (Attention le fichier sera enregistré et verrouillé dans le dossier courant, il faut retirer manuellement la protection)
SENSOR SELF NOISE |  |  |  | Decision : considering the specifications given by the manufacturer, this sensor is
CONFORM
VISA
|  |  |  | Adjusted by | A.HUE
|  |  |  | on | 2016-03-15
|  |  |  | Approved by | A.HUE
|  |  |  | Qualification | Optronic Engineer
Self Noise at 1 HZ | 5.6951e-05 | Pa/√Hz |  | on | 2016-03-17 | -84.89 | (dBPa/√Hz)
| < 0,13 | mPa/√Hz
OK
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which define the verification report. It can be used to demonstrate the link between the measure equipment and the standard national or international equipment, provided it meets the recommendations of the documentation booklet X 07-015
[8 empty] | Client | Adresse Client
[8 empty] | CEA | C.E.A. Dam. Bruyères Domaine du Grand Rué 91297 ARPAJON CEDEX FRANCE
[8 empty] | CTBTO | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
[8 empty] | IPGP | Université Paris Diderot Paris 7 Bat Lamarck UFR STEP 35 rue H.Brion 6ème étage 75205 Paris CEDEX 13
[8 empty] | SANDIA | C/O SANDIATIONAL LABS ALBUQUERQUE, NM 87123 UNITED STATES
[8 empty] | ISLA | Haiwaï Institute of Geophysics & Planetology 73-984 Makako Bay Drive KAILUA-KONA HAWAÏ
ENREGISTREMENT DEFINITIF
DEMARRER LA SAISIE

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.0522 | 11.703
0.097656 | -33.996 | 7.1128
0.12207 | -33.9114 | 5.3675
0.14648 | -33.9316 | 3.8917
0.1709 | -33.9118 | 2.9134
0.19531 | -33.9218 | 2.5981
0.21973 | -33.9867 | 2.6991
0.24414 | -33.9421 | 1.7484
0.26855 | -33.9319 | 1.9574
0.29297 | -33.9268 | 1.0598
0.31738 | -33.9486 | 1.0965
0.3418 | -33.9185 | 1.1952
0.36621 | -33.9978 | 0.91687
0.39062 | -33.9354 | 0.88913
0.41504 | -33.9903 | 0.12041
0.43945 | -33.9812 | 0.14023
0.46387 | -33.9703 | 0.003484
0.48828 | -33.9269 | 0.17776
0.5127 | -33.9815 | -0.13354
0.53711 | -33.9392 | -0.35007
0.56152 | -33.9967 | -0.22117
0.58594 | -33.9911 | -0.30432
0.61035 | -33.9882 | -0.25964
0.63477 | -33.9698 | -0.10794
0.65918 | -34.0065 | -0.64541
0.68359 | -33.9898 | -0.35656
0.70801 | -33.9584 | -1.0525
0.73242 | -33.975 | -0.61484
0.75684 | -33.9891 | -1.0053
0.78125 | -34.0181 | -0.43034
0.80566 | -33.9977 | -0.87915
0.83008 | -33.9686 | -1.0288
0.85449 | -33.9795 | -1.1366
0.87891 | -33.9739 | -1.0403
0.90332 | -33.9347 | -1.331
0.92773 | -33.9909 | -0.6514
0.95215 | -34.0197 | -1.1931
0.97656 | -34.0407 | -1.2863
1.001 | -34.0177 | -1.2699
1.0254 | -34.0596 | -1.5145
1.0498 | -34.0291 | -1.6956
1.0742 | -34.0178 | -1.7097
1.0986 | -33.9955 | -1.3534
1.123 | -33.9676 | -1.6383
1.1475 | -34.0345 | -1.4125
1.1719 | -34.0268 | -1.7785
1.1963 | -34.0097 | -1.7181
1.2207 | -34.0349 | -1.8515
1.2451 | -34.0173 | -1.8685
1.2695 | -33.9454 | -2.132
1.2939 | -33.9725 | -2.3875
1.3184 | -33.9679 | -1.8326
1.3428 | -34.0706 | -1.9638
1.3672 | -34.0296 | -2.1233
1.3916 | -34.0119 | -2.3383
1.416 | -34.0081 | -2.1395
1.4404 | -34.0063 | -2.2833
1.4648 | -33.9313 | -2.1465
1.4893 | -33.9765 | -2.4928
1.5137 | -33.9917 | -2.2967
1.5381 | -34.0025 | -2.3394
1.5625 | -33.9837 | -2.5587
1.5869 | -34.0301 | -2.615
1.6113 | -34.0762 | -2.4823
1.6357 | -34.0164 | -2.8905
1.6602 | -34.0148 | -2.4935
1.6846 | -33.9269 | -2.8118
1.709 | -33.9794 | -2.8684
1.7334 | -34.0358 | -2.561
1.7578 | -34.0243 | -2.9106
1.7822 | -33.9615 | -3.0313
1.8066 | -34.0006 | -2.8718
1.8311 | -33.9891 | -3.2462
1.8555 | -34.08 | -3.2267
1.8799 | -34.0712 | -3.1998
1.9043 | -34.0219 | -3.1519
1.9287 | -34.0031 | -3.4287
1.9531 | -34.0658 | -3.2989
1.9775 | -33.9987 | -2.8153
2.002 | -33.9859 | -3.287
2.0264 | -33.9757 | -3.5596
2.0508 | -34.0601 | -3.5645
2.0752 | -34.0289 | -3.1304
2.0996 | -34.0903 | -3.6925
2.124 | -34.0746 | -3.8649
2.1484 | -34.0238 | -3.8043
2.1729 | -34.0113 | -3.673
2.1973 | -33.9714 | -3.8408
2.2217 | -33.9801 | -3.4671
2.2461 | -34.0375 | -3.7954
2.2705 | -34.0503 | -3.9238
2.2949 | -34.0253 | -3.766
2.3193 | -34.0409 | -4.1209
2.3438 | -34.0217 | -3.9308
2.3682 | -34.0678 | -4.1542
2.3926 | -34.0019 | -3.9193
2.417 | -34.0665 | -4.4155
2.4414 | -34.0212 | -3.545
2.4658 | -34.0273 | -4.2773
2.4902 | -34.0531 | -4.3144
2.5146 | -34.0189 | -3.8046
2.5391 | -34.0546 | -4.1155
2.5635 | -33.9994 | -4.5954
2.5879 | -34.0639 | -4.2686
2.6123 | -34.0479 | -4.4172
2.6367 | -34.0401 | -4.6659
2.6611 | -34.0647 | -4.1701
2.6855 | -34.017 | -4.4707
2.71 | -34.0317 | -5.024
2.7344 | -34.0276 | -4.5223
2.7588 | -34.0573 | -4.5848
2.7832 | -34.0872 | -4.8102
2.8076 | -34.022 | -4.4888
2.832 | -34.0554 | -4.9982
2.8564 | -34.074 | -4.8553
2.8809 | -34.0697 | -5.2606
2.9053 | -34.0742 | -4.6962
2.9297 | -34.0441 | -5.0659
2.9541 | -34.0839 | -5.1767
2.9785 | -33.97 | -5.2114
3.0029 | -34.0266 | -5.0401
3.0273 | -34.0824 | -5.3663
3.0518 | -34.0469 | -5.3559
3.0762 | -34.0027 | -5.3681
3.1006 | -34.0619 | -5.1166
3.125 | -34.0841 | -5.6533
3.1494 | -33.9751 | -5.8945
3.1738 | -34.0361 | -5.5672
3.1982 | -34.073 | -6.1516
3.2227 | -34.0334 | -5.0844
3.2471 | -34.062 | -5.8284
3.2715 | -34.0493 | -4.9275
3.2959 | -34.0331 | -5.1324
3.3203 | -34.0262 | -5.2159
3.3447 | -34.0654 | -5.45
3.3691 | -34.0751 | -5.797
3.3936 | -34.0364 | -5.9755
3.418 | -34.0887 | -5.79
3.4424 | -34.0922 | -5.9571
3.4668 | -34.1095 | -5.632
3.4912 | -34.0624 | -5.7961
3.5156 | -34.1316 | -5.8216
3.54 | -34.0413 | -5.8302
3.5645 | -34.0181 | -6.1761
3.5889 | -34.0589 | -6.3781
3.6133 | -34.08 | -6.1443
3.6377 | -34.0732 | -6.2586
3.6621 | -34.0595 | -6.5587
3.6865 | -34.0859 | -6.1473
3.7109 | -34.132 | -6.4515
3.7354 | -34.0379 | -6.2266
3.7598 | -34.0729 | -6.7211
3.7842 | -34.0723 | -6.7547
3.8086 | -34.0675 | -6.4601
3.833 | -34.0854 | -6.3394
3.8574 | -34.0899 | -6.8546
3.8818 | -34.1081 | -6.7856
3.9062 | -34.1099 | -6.4652
3.9307 | -34.0945 | -6.4954
3.9551 | -34.095 | -6.8822
3.9795 | -34.1583 | -6.4146
4.0039 | -34.082 | -6.7287
4.0283 | -34.0841 | -7.2348
4.0527 | -34.0511 | -6.763
4.0771 | -34.1239 | -6.9839
4.1016 | -34.13 | -6.9422
4.126 | -34.0826 | -7.4513
4.1504 | -34.0536 | -7.2787
4.1748 | -34.1203 | -7.2836
4.1992 | -34.0963 | -7.2478
4.2236 | -34.071 | -7.756
4.248 | -34.0955 | -6.6871
4.2725 | -34.1296 | -7.3236
4.2969 | -34.1138 | -7.0943
4.3213 | -34.1033 | -7.2033
4.3457 | -34.085 | -7.6093
4.3701 | -34.0916 | -7.6353
4.3945 | -34.0851 | -7.2632
4.4189 | -34.1379 | -7.6868
4.4434 | -34.0501 | -7.4339
4.4678 | -34.131 | -7.6306
4.4922 | -34.1776 | -7.6808
4.5166 | -34.0926 | -7.8226
4.541 | -34.1067 | -7.8466
4.5654 | -34.0954 | -7.8546
4.5898 | -34.1709 | -8.0271
4.6143 | -34.0132 | -7.9313
4.6387 | -34.1287 | -7.9066
4.6631 | -34.1122 | -7.9162
4.6875 | -34.1503 | -7.9333
4.7119 | -34.0836 | -7.6217
4.7363 | -34.1473 | -7.9908
4.7607 | -34.1132 | -7.6216
4.7852 | -34.0376 | -8.8144
4.8096 | -34.1538 | -7.921
4.834 | -34.0826 | -8.5966
4.8584 | -34.1263 | -8.2352
4.8828 | -34.1372 | -8.3888
4.9072 | -34.1648 | -8.5489
4.9316 | -34.1545 | -8.635
4.9561 | -34.1677 | -8.454
4.9805 | -34.1098 | -8.3231
5.0049 | -34.1244 | -8.6928
5.0293 | -34.1042 | -8.8005
5.0537 | -34.1808 | -7.9692
5.0781 | -34.0813 | -8.3355
5.1025 | -34.1533 | -8.4009
5.127 | -34.1747 | -8.6927
5.1514 | -34.0786 | -8.839
5.1758 | -34.1261 | -8.4042
5.2002 | -34.1157 | -9.0827
5.2246 | -34.1131 | -8.8358
5.249 | -34.2141 | -9.0477
5.2734 | -34.1233 | -8.2172
5.2979 | -34.1468 | -9.1795
5.3223 | -34.1607 | -9.0365
5.3467 | -34.1835 | -8.8339
5.3711 | -34.1166 | -9.3106
5.3955 | -34.103 | -9.3768
5.4199 | -34.144 | -9.3935
5.4443 | -34.1774 | -9.5423
5.4688 | -34.2006 | -9.4527
5.4932 | -34.1222 | -9.1443
5.5176 | -34.1645 | -9.5004
5.542 | -34.1797 | -9.4801
5.5664 | -34.1677 | -9.0955
5.5908 | -34.1799 | -9.3292
5.6152 | -34.1992 | -9.9513
5.6396 | -34.1547 | -9.432
5.6641 | -34.0934 | -9.4605
5.6885 | -34.2013 | -9.2231
5.7129 | -34.1079 | -9.8307
5.7373 | -34.1532 | -9.5062
5.7617 | -34.0994 | -9.6445
5.7861 | -34.1779 | -9.1974
5.8105 | -34.0988 | -9.8337
5.835 | -34.1835 | -9.7473
5.8594 | -34.2052 | -10.1675
5.8838 | -34.1505 | -10.1152
5.9082 | -34.1317 | -9.8977
5.9326 | -34.2033 | -10.3055
5.957 | -34.1148 | -9.6286
5.9814 | -34.1545 | -10.2535
6.0059 | -34.1301 | -9.9988
6.0303 | -34.1606 | -10.3081
6.0547 | -34.1329 | -10.4639
6.0791 | -34.1136 | -10.0778
6.1035 | -34.2107 | -10.2791
6.1279 | -34.1885 | -10.5839
6.1523 | -34.2643 | -10.5404
6.1768 | -34.1319 | -10.3301
6.2012 | -34.1825 | -10.3041
6.2256 | -34.1884 | -10.7686
6.25 | -34.1598 | -10.3247
6.2744 | -34.165 | -10.7637
6.2988 | -34.2421 | -10.5255
6.3232 | -34.1061 | -10.4042
6.3477 | -34.2405 | -11.1974
6.3721 | -34.1777 | -10.6442
6.3965 | -34.1895 | -11.063
6.4209 | -34.1767 | -10.6707
6.4453 | -34.1642 | -11.0063
6.4697 | -34.2553 | -10.8649
6.4941 | -34.2707 | -11.1373
6.5186 | -34.2025 | -10.9063
6.543 | -34.2342 | -11.417
6.5674 | -34.2832 | -11.5057
6.5918 | -34.1956 | -11.4947
6.6162 | -34.2121 | -10.6547
6.6406 | -34.1523 | -11.3761
6.665 | -34.1844 | -10.7979
6.6895 | -34.1379 | -11.2011
6.7139 | -34.2069 | -11.3706
6.7383 | -34.1412 | -11.6306
6.7627 | -34.2008 | -11.2647
6.7871 | -34.2627 | -11.4457
6.8115 | -34.2758 | -11.6668
6.8359 | -34.2257 | -11.5615
6.8604 | -34.2336 | -11.7471
6.8848 | -34.2364 | -11.1835
6.9092 | -34.2303 | -11.646
6.9336 | -34.2505 | -11.4605
6.958 | -34.2387 | -12.1193
6.9824 | -34.1684 | -11.4043
7.0068 | -34.1721 | -12.0057
7.0312 | -34.2213 | -12.0633
7.0557 | -34.2208 | -11.6479
7.0801 | -34.3015 | -11.6625
7.1045 | -34.2489 | -11.9452
7.1289 | -34.2813 | -12.0194
7.1533 | -34.2447 | -11.6081
7.1777 | -34.2585 | -11.809
7.2021 | -34.2426 | -11.7475
7.2266 | -34.2888 | -12.1885
7.251 | -34.2244 | -12.1352
7.2754 | -34.2299 | -12.4408
7.2998 | -34.2046 | -12.002
7.3242 | -34.3043 | -12.3979
7.3486 | -34.261 | -11.8976
7.373 | -34.263 | -12.3472
7.3975 | -34.1749 | -12.1438
7.4219 | -34.2615 | -12.3918
7.4463 | -34.2398 | -12.3705
7.4707 | -34.2115 | -12.5948
7.4951 | -34.3054 | -12.6463
7.5195 | -34.1828 | -12.6428
7.5439 | -34.2707 | -12.0411
7.5684 | -34.2831 | -12.5773
7.5928 | -34.2815 | -12.4245
7.6172 | -34.2713 | -12.659
7.6416 | -34.2228 | -12.9837
7.666 | -34.318 | -13.0174
7.6904 | -34.3085 | -12.4226
7.7148 | -34.2143 | -12.9134
7.7393 | -34.3566 | -12.9845
7.7637 | -34.297 | -13.0434
7.7881 | -34.2748 | -13.0442
7.8125 | -34.3054 | -13.0629
7.8369 | -34.2401 | -13.2243
7.8613 | -34.2628 | -13.017
7.8857 | -34.3036 | -13.2592
7.9102 | -34.2405 | -13.0163
7.9346 | -34.2946 | -13.1128
7.959 | -34.2858 | -12.8612
7.9834 | -34.281 | -13.3682
8.0078 | -34.3094 | -13.3719
8.0322 | -34.3058 | -13.4099
8.0566 | -34.3215 | -12.9475
8.0811 | -34.2628 | -13.4245
8.1055 | -34.335 | -13.8322
8.1299 | -34.2729 | -13.5446
8.1543 | -34.3211 | -13.5898
8.1787 | -34.2755 | -13.7784
8.2031 | -34.385 | -13.6505
8.2275 | -34.2527 | -13.4269
8.252 | -34.3205 | -13.8586
8.2764 | -34.3028 | -13.3822
8.3008 | -34.2511 | -13.7063
8.3252 | -34.2887 | -14.1184
8.3496 | -34.3579 | -13.9491
8.374 | -34.288 | -14.1869
8.3984 | -34.2868 | -14.0529
8.4229 | -34.3584 | -13.5875
8.4473 | -34.2686 | -13.972
8.4717 | -34.2954 | -13.8902
8.4961 | -34.2517 | -13.822
8.5205 | -34.3265 | -14.1394
8.5449 | -34.2646 | -13.8757
8.5693 | -34.3767 | -14.1638
8.5938 | -34.2512 | -13.9861
8.6182 | -34.3431 | -14.5862
8.6426 | -34.3346 | -14.0119
8.667 | -34.3207 | -14.2591
8.6914 | -34.2842 | -14.1468
8.7158 | -34.314 | -14.5057
8.7402 | -34.3132 | -14.1068
8.7646 | -34.3516 | -14.7562
8.7891 | -34.3623 | -14.3364
8.8135 | -34.3197 | -14.2056
8.8379 | -34.3182 | -14.281
8.8623 | -34.329 | -14.8678
8.8867 | -34.3817 | -14.5012
8.9111 | -34.394 | -14.3406
8.9355 | -34.4061 | -14.3288
8.96 | -34.3344 | -14.2644
8.9844 | -34.3931 | -14.8087
9.0088 | -34.3373 | -14.7968
9.0332 | -34.3561 | -14.9057
9.0576 | -34.3226 | -14.758
9.082 | -34.3717 | -15.2376
9.1064 | -34.3278 | -14.7966
9.1309 | -34.4226 | -15.1589
9.1553 | -34.3308 | -15.1395
9.1797 | -34.3566 | -15.1915
9.2041 | -34.4348 | -14.8906
9.2285 | -34.367 | -15.4578
9.2529 | -34.3458 | -14.9414
9.2773 | -34.3693 | -15.2264
9.3018 | -34.3548 | -14.9491
9.3262 | -34.3455 | -15.346
9.3506 | -34.4328 | -15.4136
9.375 | -34.3198 | -15.1991
9.3994 | -34.3437 | -15.436
9.4238 | -34.3613 | -15.3822
9.4482 | -34.4125 | -15.4708
9.4727 | -34.431 | -15.4222
9.4971 | -34.423 | -15.8549
9.5215 | -34.3534 | -15.4639
9.5459 | -34.4649 | -15.8395
9.5703 | -34.3502 | -15.5547
9.5947 | -34.4281 | -15.7286
9.6191 | -34.3471 | -15.4701
9.6436 | -34.3836 | -15.9852
9.668 | -34.4552 | -15.6146
9.6924 | -34.4365 | -16.4187
9.7168 | -34.4216 | -16.5755
9.7412 | -34.4083 | -15.5778
9.7656 | -34.3677 | -15.9127
9.79 | -34.4694 | -15.5787
9.8145 | -34.4229 | -15.9372
9.8389 | -34.3361 | -16.1076
9.8633 | -34.4027 | -16.3857
9.8877 | -34.3767 | -15.8554
9.9121 | -34.4033 | -16.5408
9.9365 | -34.3804 | -16.2909
9.9609 | -34.4621 | -16.2681
9.9854 | -34.3702 | -15.9727
10.0098 | -34.4028 | -16.129

## ACOUS_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.2266 | 9.7896
0.097656 | -34.222 | 4.5376
0.12207 | -33.983 | 3.7611
0.14648 | -34.1005 | 1.3039
0.1709 | -34.0117 | 1.206
0.19531 | -34.0238 | 0.23616
0.21973 | -34.021 | 0.25762
0.24414 | -34.0359 | -1.1051
0.26855 | -34.0753 | -0.26247
0.29297 | -33.9858 | -1.5846
0.31738 | -34.0429 | -1.1998
0.3418 | -33.9954 | -1.7795
0.36621 | -34.0638 | -1.3809
0.39062 | -33.9585 | -2.1943
0.41504 | -34.0841 | -1.746
0.43945 | -33.9806 | -2.1139
0.46387 | -34.0624 | -2.393
0.48828 | -33.9423 | -2.3292
0.5127 | -34.0917 | -2.7145
0.53711 | -33.9692 | -2.3905
0.56152 | -34.0763 | -3.1429
0.58594 | -33.9929 | -2.5494
0.61035 | -34.0645 | -3.3316
0.63477 | -34.0127 | -2.6837
0.65918 | -34.0565 | -3.5689
0.68359 | -34.0139 | -2.7707
0.70801 | -34.0052 | -3.6556
0.73242 | -34.052 | -3.0424
0.75684 | -34.014 | -3.8963
0.78125 | -34.0667 | -3.22
0.80566 | -33.9865 | -4.0342
0.83008 | -34.0995 | -3.6111
0.85449 | -33.9672 | -4.0551
0.87891 | -34.0573 | -3.8003
0.90332 | -33.9624 | -3.8731
0.92773 | -34.1025 | -3.8983
0.95215 | -33.996 | -3.8703
0.97656 | -34.105 | -4.4125
1.001 | -33.9964 | -4.0069
1.0254 | -34.0952 | -4.5709
1.0498 | -34.0349 | -4.0829
1.0742 | -34.0965 | -5.0363
1.0986 | -34.0207 | -4.261
1.123 | -34.0393 | -5.0218
1.1475 | -34.0462 | -4.2874
1.1719 | -34.0432 | -5.1276
1.1963 | -34.075 | -4.2447
1.2207 | -34.0245 | -5.2847
1.2451 | -34.1135 | -4.4883
1.2695 | -34.0077 | -5.1656
1.2939 | -34.1075 | -4.6808
1.3184 | -33.9799 | -5.1138
1.3428 | -34.1391 | -4.8653
1.3672 | -34.0173 | -5.0978
1.3916 | -34.1223 | -5.2818
1.416 | -33.9842 | -5.1776
1.4404 | -34.1125 | -5.4048
1.4648 | -33.9993 | -4.9029
1.4893 | -34.1169 | -5.4406
1.5137 | -34.0439 | -4.9277
1.5381 | -34.1113 | -5.8759
1.5625 | -34.0699 | -5.0747
1.5869 | -34.0963 | -5.9666
1.6113 | -34.0941 | -5.2999
1.6357 | -34.0413 | -6.2905
1.6602 | -34.086 | -5.486
1.6846 | -34.0304 | -6.0861
1.709 | -34.1148 | -5.5409
1.7334 | -33.9967 | -6.1409
1.7578 | -34.1392 | -5.8338
1.7822 | -33.9936 | -6.0035
1.8066 | -34.1321 | -5.9449
1.8311 | -34.0018 | -5.9466
1.8555 | -34.1613 | -6.1503
1.8799 | -34.0284 | -5.9839
1.9043 | -34.1236 | -6.5809
1.9287 | -34.005 | -6.0344
1.9531 | -34.1034 | -6.715
1.9775 | -34.0409 | -5.8135
2.002 | -34.096 | -6.8184
2.0264 | -34.0603 | -6.0052
2.0508 | -34.0705 | -6.7536
2.0752 | -34.1043 | -6.0981
2.0996 | -34.0797 | -6.9544
2.124 | -34.1312 | -6.275
2.1484 | -34.0554 | -7.2157
2.1729 | -34.1337 | -6.6155
2.1973 | -34.0399 | -7.003
2.2217 | -34.143 | -6.7872
2.2461 | -34.0416 | -7.0075
2.2705 | -34.1861 | -6.9786
2.2949 | -34.0525 | -7.0667
2.3193 | -34.1542 | -7.2283
2.3438 | -34.0527 | -6.8874
2.3682 | -34.1209 | -7.4473
2.3926 | -34.0443 | -6.766
2.417 | -34.1148 | -7.5466
2.4414 | -34.0843 | -6.7248
2.4658 | -34.1409 | -7.6281
2.4902 | -34.1273 | -6.8623
2.5146 | -34.1071 | -7.8909
2.5391 | -34.1152 | -7.1418
2.5635 | -34.0861 | -7.8378
2.5879 | -34.1434 | -7.2395
2.6123 | -34.0586 | -7.7316
2.6367 | -34.1516 | -7.3814
2.6611 | -34.0577 | -7.8554
2.6855 | -34.1792 | -7.5445
2.71 | -34.0364 | -7.8363
2.7344 | -34.1837 | -7.8249
2.7588 | -34.0717 | -7.5461
2.7832 | -34.1566 | -8.1484
2.8076 | -34.0532 | -7.6558
2.832 | -34.1682 | -8.1699
2.8564 | -34.1083 | -7.7561
2.8809 | -34.1405 | -8.6112
2.9053 | -34.0939 | -7.7797
2.9297 | -34.1259 | -8.6205
2.9541 | -34.1387 | -7.9252
2.9785 | -34.0891 | -8.6997
3.0029 | -34.1472 | -7.971
3.0273 | -34.0907 | -8.6851
3.0518 | -34.1841 | -8.0698
3.0762 | -34.0868 | -8.6126
3.1006 | -34.1836 | -8.2504
3.125 | -34.0719 | -8.6241
3.1494 | -34.1897 | -8.5247
3.1738 | -34.0738 | -8.7403
3.1982 | -34.1903 | -8.8536
3.2227 | -34.0538 | -8.5723
3.2471 | -34.1688 | -8.9799
3.2715 | -34.0979 | -8.4996
3.2959 | -34.1748 | -9.0035
3.3203 | -34.1053 | -8.4413
3.3447 | -34.1472 | -9.2815
3.3691 | -34.1343 | -8.5563
3.3936 | -34.1474 | -9.3052
3.418 | -34.155 | -8.6365
3.4424 | -34.1235 | -9.5245
3.4668 | -34.1914 | -8.7835
3.4912 | -34.1149 | -9.6723
3.5156 | -34.2237 | -9.0533
3.54 | -34.0735 | -9.5176
3.5645 | -34.2074 | -9.2926
3.5889 | -34.1127 | -9.4697
3.6133 | -34.2176 | -9.3964
3.6377 | -34.0801 | -9.5065
3.6621 | -34.2081 | -9.6922
3.6865 | -34.0996 | -9.2352
3.7109 | -34.2172 | -9.6929
3.7354 | -34.1359 | -9.3924
3.7598 | -34.2046 | -9.959
3.7842 | -34.1372 | -9.5147
3.8086 | -34.1589 | -10.1317
3.833 | -34.1457 | -9.4111
3.8574 | -34.159 | -10.0962
3.8818 | -34.1997 | -9.4718
3.9062 | -34.1338 | -10.3393
3.9307 | -34.1881 | -9.5216
3.9551 | -34.126 | -10.2048
3.9795 | -34.2526 | -9.9026
4.0039 | -34.1125 | -10.2157
4.0283 | -34.2286 | -10.0672
4.0527 | -34.089 | -10.2675
4.0771 | -34.2372 | -10.3436
4.1016 | -34.1416 | -10.1717
4.126 | -34.24 | -10.5227
4.1504 | -34.1367 | -10.2064
4.1748 | -34.2285 | -10.8503
4.1992 | -34.1469 | -10.2855
4.2236 | -34.1859 | -10.9874
4.248 | -34.144 | -10.2024
4.2725 | -34.1654 | -10.946
4.2969 | -34.1691 | -10.1195
4.3213 | -34.1846 | -11.0448
4.3457 | -34.1916 | -10.368
4.3701 | -34.1467 | -11.1536
4.3945 | -34.2209 | -10.482
4.4189 | -34.1339 | -11.0522
4.4434 | -34.2643 | -10.5966
4.4678 | -34.1608 | -11.1179
4.4922 | -34.2503 | -11.097
4.5166 | -34.106 | -11.082
4.541 | -34.2542 | -11.1826
4.5654 | -34.134 | -10.9813
4.5898 | -34.2504 | -11.4614
4.6143 | -34.1407 | -10.8336
4.6387 | -34.2298 | -11.6497
4.6631 | -34.1748 | -10.9232
4.6875 | -34.2181 | -11.7974
4.7119 | -34.1859 | -11.0167
4.7363 | -34.1985 | -11.5157
4.7607 | -34.2084 | -11.0284
4.7852 | -34.2027 | -11.9232
4.8096 | -34.2414 | -11.2664
4.834 | -34.1636 | -11.9583
4.8584 | -34.2792 | -11.2191
4.8828 | -34.1533 | -11.6767
4.9072 | -34.2855 | -11.4956
4.9316 | -34.1545 | -11.9072
4.9561 | -34.2788 | -11.8769
4.9805 | -34.148 | -11.6811
5.0049 | -34.2566 | -12.1407
5.0293 | -34.1192 | -11.7133
5.0537 | -34.2417 | -12.2672
5.0781 | -34.1588 | -11.5263
5.1025 | -34.2613 | -12.4353
5.127 | -34.1854 | -11.7215
5.1514 | -34.2501 | -12.5425
5.1758 | -34.2223 | -11.7883
5.2002 | -34.2314 | -12.6365
5.2246 | -34.2511 | -11.8232
5.249 | -34.1707 | -12.5777
5.2734 | -34.2518 | -11.8676
5.2979 | -34.1943 | -12.7279
5.3223 | -34.3148 | -12.2393
5.3467 | -34.1675 | -12.5433
5.3711 | -34.3009 | -12.4053
5.3955 | -34.1505 | -12.6384
5.4199 | -34.3177 | -12.726
5.4443 | -34.1826 | -12.554
5.4688 | -34.2777 | -13.057
5.4932 | -34.1612 | -12.4943
5.5176 | -34.2716 | -13.0005
5.542 | -34.1935 | -12.5619
5.5664 | -34.2706 | -13.2618
5.5908 | -34.2085 | -12.6122
5.6152 | -34.2446 | -13.4507
5.6396 | -34.2228 | -12.8241
5.6641 | -34.2157 | -13.4923
5.6885 | -34.2599 | -12.721
5.7129 | -34.1914 | -13.5048
5.7373 | -34.2695 | -12.8669
5.7617 | -34.1984 | -13.3806
5.7861 | -34.3343 | -12.7706
5.8105 | -34.1823 | -13.4193
5.835 | -34.3473 | -13.2456
5.8594 | -34.1817 | -13.3633
5.8838 | -34.327 | -13.455
5.9082 | -34.1988 | -13.409
5.9326 | -34.3162 | -13.8271
5.957 | -34.2171 | -13.107
5.9814 | -34.3101 | -13.863
6.0059 | -34.2297 | -13.3275
6.0303 | -34.2784 | -14.0489
6.0547 | -34.2332 | -13.3603
6.0791 | -34.2468 | -13.9366
6.1035 | -34.2645 | -13.1257
6.1279 | -34.239 | -14.2894
6.1523 | -34.2918 | -13.5415
6.1768 | -34.2085 | -14.3599
6.2012 | -34.3016 | -13.7776
6.2256 | -34.2098 | -14.0781
6.25 | -34.3391 | -13.7733
6.2744 | -34.2212 | -13.9948
6.2988 | -34.3207 | -14.1513
6.3232 | -34.22 | -13.921
6.3477 | -34.3335 | -14.0164
6.3721 | -34.2022 | -13.8601
6.3965 | -34.3252 | -14.4108
6.4209 | -34.2255 | -13.9821
6.4453 | -34.3062 | -14.5382
6.4697 | -34.256 | -13.8979
6.4941 | -34.3027 | -14.9293
6.5186 | -34.2971 | -14.0175
6.543 | -34.2985 | -15.0334
6.5674 | -34.3276 | -14.1146
6.5918 | -34.269 | -15.1168
6.6162 | -34.3136 | -14.3894
6.6406 | -34.2126 | -14.9489
6.665 | -34.3593 | -14.4727
6.6895 | -34.2296 | -14.731
6.7139 | -34.3636 | -14.7757
6.7383 | -34.224 | -14.884
6.7627 | -34.3351 | -14.945
6.7871 | -34.1954 | -14.7409
6.8115 | -34.3773 | -15.0486
6.8359 | -34.2672 | -14.9858
6.8604 | -34.3505 | -15.2885
6.8848 | -34.2367 | -14.7496
6.9092 | -34.33 | -15.4555
6.9336 | -34.3173 | -14.8976
6.958 | -34.3101 | -15.6593
6.9824 | -34.3154 | -14.8282
7.0068 | -34.3047 | -15.6295
7.0312 | -34.3313 | -14.8412
7.0557 | -34.2954 | -15.5478
7.0801 | -34.3539 | -14.923
7.1045 | -34.2988 | -15.4723
7.1289 | -34.3985 | -15.2123
7.1533 | -34.2585 | -15.6052
7.1777 | -34.3882 | -15.4051
7.2021 | -34.2797 | -15.6725
7.2266 | -34.4148 | -15.6169
7.251 | -34.315 | -15.6619
7.2754 | -34.3831 | -15.7999
7.2998 | -34.2586 | -15.2831
7.3242 | -34.3522 | -16.049
7.3486 | -34.3065 | -15.3774
7.373 | -34.3526 | -16.1934
7.3975 | -34.3352 | -15.3701
7.4219 | -34.3539 | -16.371
7.4463 | -34.3385 | -15.5529
7.4707 | -34.3075 | -16.3102
7.4951 | -34.3251 | -15.7302
7.5195 | -34.3052 | -16.2524
7.5439 | -34.378 | -15.8642
7.5684 | -34.2783 | -16.178
7.5928 | -34.3615 | -16.01
7.6172 | -34.2563 | -16.1971
7.6416 | -34.4179 | -16.1813
7.666 | -34.2738 | -15.9752
7.6904 | -34.4411 | -16.4898
7.7148 | -34.3212 | -16.179
7.7393 | -34.4059 | -16.6764
7.7637 | -34.3625 | -16.1997
7.7881 | -34.4041 | -16.9902
7.8125 | -34.3451 | -16.2868
7.8369 | -34.348 | -17.051
7.8613 | -34.356 | -16.1237
7.8857 | -34.3471 | -17.2282
7.9102 | -34.4021 | -16.3001
7.9346 | -34.3069 | -17.1386
7.959 | -34.412 | -16.4936
7.9834 | -34.3214 | -17.1448
8.0078 | -34.4334 | -16.6353
8.0322 | -34.3295 | -17.1393
8.0566 | -34.4494 | -16.9178
8.0811 | -34.3222 | -16.8425
8.1055 | -34.4373 | -16.8847
8.1299 | -34.2963 | -16.8684
8.1543 | -34.4196 | -17.3247
8.1787 | -34.3266 | -16.8137
8.2031 | -34.3823 | -17.2406
8.2275 | -34.3584 | -16.8468
8.252 | -34.398 | -17.5678
8.2764 | -34.3841 | -17.0437
8.3008 | -34.397 | -17.6341
8.3252 | -34.4008 | -17.0946
8.3496 | -34.3853 | -17.8267
8.374 | -34.4423 | -16.982
8.3984 | -34.3562 | -18.0327
8.4229 | -34.4219 | -17.1998
8.4473 | -34.3651 | -17.824
8.4717 | -34.4361 | -17.2937
8.4961 | -34.3269 | -17.933
8.5205 | -34.4636 | -17.3105
8.5449 | -34.3233 | -17.8072
8.5693 | -34.4925 | -17.7284
8.5938 | -34.3373 | -17.4218
8.6182 | -34.476 | -18.162
8.6426 | -34.3593 | -17.4856
8.667 | -34.4317 | -18.2332
8.6914 | -34.3664 | -17.2679
8.7158 | -34.409 | -18.4538
8.7402 | -34.3769 | -17.5465
8.7646 | -34.4047 | -18.3127
8.7891 | -34.4445 | -17.7381
8.8135 | -34.3982 | -18.6667
8.8379 | -34.4779 | -17.6218
8.8623 | -34.3888 | -18.299
8.8867 | -34.4947 | -18.136
8.9111 | -34.3677 | -18.2609
8.9355 | -34.4851 | -18.3522
8.96 | -34.374 | -18.5468
8.9844 | -34.4922 | -18.5246
9.0088 | -34.3772 | -18.4593
9.0332 | -34.4943 | -18.5598
9.0576 | -34.3381 | -18.3143
9.082 | -34.5105 | -18.7879
9.1064 | -34.4069 | -18.0303
9.1309 | -34.4781 | -18.9584
9.1553 | -34.4 | -18.2642
9.1797 | -34.4405 | -19.1315
9.2041 | -34.4358 | -18.4199
9.2285 | -34.4416 | -19.1263
9.2529 | -34.4222 | -18.2083
9.2773 | -34.3867 | -19.3835
9.3018 | -34.4668 | -18.4114
9.3262 | -34.4246 | -19.1061
9.3506 | -34.5281 | -18.7063
9.375 | -34.3842 | -18.9431
9.3994 | -34.5042 | -18.6896
9.4238 | -34.4118 | -18.9546
9.4482 | -34.5443 | -18.8031
9.4727 | -34.404 | -19.0256
9.4971 | -34.5049 | -19.39
9.5215 | -34.3892 | -18.9237
9.5459 | -34.4877 | -19.621
9.5703 | -34.43 | -18.7721
9.5947 | -34.4683 | -19.7535
9.6191 | -34.4343 | -19.0494
9.6436 | -34.4948 | -19.743
9.668 | -34.475 | -18.9921
9.6924 | -34.4096 | -19.7442
9.7168 | -34.4849 | -19.1269
9.7412 | -34.4385 | -19.8693
9.7656 | -34.5235 | -19.0777
9.79 | -34.3978 | -19.8455
9.8145 | -34.5333 | -19.2
9.8389 | -34.4061 | -20.1992
9.8633 | -34.5421 | -19.562
9.8877 | -34.3927 | -19.7181
9.9121 | -34.5272 | -19.6544
9.9365 | -34.413 | -19.8381
9.9609 | -34.5531 | -19.8645
9.9854 | -34.4386 | -19.6663
10.0098 | -34.5655 | -20.4018

## SEALED_NOISE
dBPa/√Hz | Frequency
0.009572 | -29.5267
0.009881 | -29.4016
0.010189 | -31.4653
0.010498 | -34.2467
0.010807 | -36.4938
0.011116 | -38.2953
0.011424 | -39.2847
0.011733 | -38.7916
0.012042 | -37.3768
0.012351 | -34.8519
0.012659 | -33.5064
0.012968 | -35.4698
0.013277 | -39.0172
0.013586 | -37.9282
0.013895 | -34.0336
0.014203 | -33.0934
0.014512 | -35.4876
0.014821 | -40.0101
0.01513 | -44.2289
0.015438 | -44.3652
0.015747 | -38.7225
0.016056 | -35.3553
0.016365 | -35.3619
0.016673 | -38.4056
0.016982 | -40.4826
0.017291 | -36.0714
0.0176 | -34.8724
0.017909 | -35.5841
0.018217 | -37.9799
0.018526 | -40.5319
0.018835 | -40.7758
0.019144 | -39.6815
0.019452 | -39.361
0.019761 | -40.3566
0.02007 | -42.2054
0.020379 | -43.4631
0.020687 | -43.5983
0.020996 | -43.5414
0.021305 | -43.1185
0.021614 | -42.2151
0.021923 | -41.6636
0.022231 | -41.7313
0.02254 | -42.1326
0.022849 | -42.0552
0.023158 | -41.2411
0.023466 | -40.5544
0.023775 | -40.0601
0.024084 | -39.7848
0.024393 | -40.3814
0.024701 | -41.9707
0.02501 | -43.1881
0.025319 | -42.4156
0.025628 | -41.1006
0.025936 | -40.7
0.026245 | -41.2731
0.026554 | -41.5168
0.026863 | -41.2452
0.027172 | -40.5225
0.02748 | -39.8291
0.027789 | -39.6153
0.028098 | -39.9478
0.028407 | -40.5668
0.028715 | -40.9888
0.029024 | -41.0357
0.029333 | -41.0635
0.029642 | -41.36
0.02995 | -41.9135
0.030259 | -42.3809
0.030568 | -42.3191
0.030877 | -41.6833
0.031186 | -40.9171
0.031494 | -40.654
0.031803 | -41.2003
0.032112 | -42.5197
0.032421 | -43.5817
0.032729 | -42.6365
0.033038 | -40.1835
0.033347 | -37.8831
0.033656 | -36.5346
0.033964 | -36.3502
0.034273 | -37.3377
0.034582 | -39.4039
0.034891 | -42.0194
0.0352 | -44.0131
0.035508 | -43.397
0.035817 | -43.5367
0.036126 | -43.3895
0.036435 | -43.5381
0.036743 | -43.9901
0.037052 | -44.5181
0.037361 | -45.0006
0.03767 | -45.2839
0.037978 | -45.6745
0.038287 | -46.0246
0.038596 | -46.4277
0.038905 | -46.7906
0.039214 | -46.9679
0.039522 | -46.7901
0.039831 | -46.3349
0.04014 | -45.5901
0.040449 | -44.9179
0.040757 | -44.5608
0.041066 | -44.7042
0.041375 | -45.2687
0.041684 | -46.045
0.041992 | -46.6087
0.042301 | -46.8159
0.04261 | -46.7566
0.042919 | -46.4872
0.043227 | -46.0218
0.043536 | -45.4854
0.043845 | -44.9095
0.044154 | -44.4712
0.044463 | -44.7166
0.044771 | -45.0276
0.04508 | -45.6584
0.045389 | -46.6011
0.045698 | -47.7739
0.046006 | -48.9019
0.046315 | -49.5032
0.046624 | -49.4136
0.046933 | -48.5512
0.047241 | -47.3225
0.04755 | -46.195
0.047859 | -45.3788
0.048168 | -44.9981
0.048477 | -45.0928
0.048785 | -45.6283
0.049094 | -46.471
0.049403 | -47.3765
0.049712 | -47.8887
0.05002 | -47.769
0.050329 | -47.279
0.050638 | -46.7527
0.050947 | -46.4815
0.051255 | -46.5836
0.051564 | -46.9687
0.051873 | -47.4883
0.052182 | -48.0287
0.052491 | -48.526
0.052799 | -49.0211
0.053108 | -49.4455
0.053417 | -49.4646
0.053726 | -49.5073
0.054034 | -49.4507
0.054343 | -49.3963
0.054652 | -49.5052
0.054961 | -49.7421
0.055269 | -50.0936
0.055578 | -50.5173
0.055887 | -50.7482
0.056196 | -50.7117
0.056505 | -50.4251
0.056813 | -49.8447
0.057122 | -49.1117
0.057431 | -48.3582
0.05774 | -47.7039
0.058048 | -47.2402
0.058357 | -46.9604
0.058666 | -46.8332
0.058975 | -46.8547
0.059283 | -46.9685
0.059592 | -47.1467
0.059901 | -47.3286
0.06021 | -47.4866
0.060518 | -47.6651
0.060827 | -47.8247
0.061136 | -48.0447
0.061445 | -48.3946
0.061754 | -48.8443
0.062062 | -49.273
0.062371 | -49.7077
0.06268 | -50.1086
0.062989 | -50.4609
0.063297 | -50.8469
0.063606 | -51.3021
0.063915 | -51.7853
0.064224 | -52.2777
0.064532 | -52.8254
0.064841 | -53.4534
0.06515 | -54.0247
0.065459 | -54.5354
0.065768 | -54.8832
0.066076 | -54.9747
0.066385 | -55.034
0.066694 | -54.979
0.067003 | -54.9532
0.067311 | -55.0626
0.06762 | -55.2582
0.067929 | -55.4634
0.068238 | -55.5988
0.068546 | -55.5865
0.068855 | -55.4876
0.069164 | -55.2673
0.069473 | -55.0192
0.069782 | -54.742
0.07009 | -54.4365
0.070399 | -54.1982
0.070708 | -53.9657
0.071017 | -53.9945
0.071325 | -53.9709
0.071634 | -54.0291
0.071943 | -54.1535
0.072252 | -54.3442
0.07256 | -54.5498
0.072869 | -54.7913
0.073178 | -55.0732
0.073487 | -55.4037
0.073795 | -55.8295
0.074104 | -56.2717
0.074413 | -56.6912
0.074722 | -57.0866
0.075031 | -57.4299
0.075339 | -57.6757
0.075648 | -57.7687
0.075957 | -57.8234
0.076266 | -57.8201
0.076574 | -57.5384
0.076883 | -57.2005
0.077192 | -56.8243
0.077501 | -56.3761
0.077809 | -55.9396
0.078118 | -55.5757
0.078427 | -55.2911
0.078736 | -55.1222
0.079045 | -55.0919
0.079353 | -55.2186
0.079662 | -55.4603
0.079971 | -55.8979
0.08028 | -56.2335
0.080588 | -56.5461
0.080897 | -56.8677
0.081206 | -57.0487
0.081515 | -57.0706
0.081823 | -56.9543
0.082132 | -56.7083
0.082441 | -56.4448
0.08275 | -56.1874
0.083059 | -55.9124
0.083367 | -55.7084
0.083676 | -55.5909
0.083985 | -55.5665
0.084294 | -55.6184
0.084602 | -55.7622
0.084911 | -55.9966
0.08522 | -56.297
0.085529 | -56.6686
0.085837 | -57.0865
0.086146 | -57.5602
0.086455 | -57.9962
0.086764 | -58.3603
0.087073 | -58.6933
0.087381 | -58.8644
0.08769 | -58.9113
0.087999 | -58.8525
0.088308 | -58.7306
0.088616 | -58.425
0.088925 | -58.193
0.089234 | -58.0126
0.089543 | -57.8945
0.089851 | -57.8319
0.09016 | -57.8121
0.090469 | -57.87
0.090778 | -57.9799
0.091086 | -58.1735
0.091395 | -58.4077
0.091704 | -58.6196
0.092013 | -58.8624
0.092322 | -59.0352
0.09263 | -59.2136
0.092939 | -59.2573
0.093248 | -59.2013
0.093557 | -59.0703
0.093865 | -58.788
0.094174 | -58.4379
0.094483 | -58.1273
0.094792 | -57.8659
0.0951 | -57.6273
0.095409 | -57.4319
0.095718 | -57.2625
0.096027 | -57.1689
0.096336 | -57.0926
0.096644 | -57.0222
0.096953 | -56.928
0.097262 | -56.8023
0.097571 | -56.6443
0.097879 | -56.5227
0.098188 | -56.4205
0.098497 | -56.3613
0.098806 | -56.3553
0.099114 | -56.4293
0.099423 | -56.5369
0.099732 | -56.6552
0.10004 | -56.7538
0.10035 | -56.9116
0.10066 | -57.0947
0.10097 | -57.2626
0.10128 | -57.384
0.10158 | -57.467
0.10189 | -57.4915
0.1022 | -57.4751
0.10251 | -57.4522
0.10282 | -57.4517
0.10313 | -57.4435
0.10344 | -57.4146
0.10375 | -57.4526
0.10405 | -57.5631
0.10436 | -57.7522
0.10467 | -57.9817
0.10498 | -58.2435
0.10529 | -58.5076
0.1056 | -58.7844
0.10591 | -59.0144
0.10622 | -59.1834
0.10652 | -59.3156
0.10683 | -59.3853
0.10714 | -59.5054
0.10745 | -59.6391
0.10776 | -59.8397
0.10807 | -60.0633
0.10838 | -60.2249
0.10869 | -60.4046
0.10899 | -60.6288
0.1093 | -60.9009
0.10961 | -61.1842
0.10992 | -61.4778
0.11023 | -61.7618
0.11054 | -62.0528
0.11085 | -62.3378
0.11116 | -62.5689
0.11146 | -62.841
0.11177 | -63.084
0.11208 | -63.293
0.11239 | -63.4672
0.1127 | -63.6671
0.11301 | -63.828
0.11332 | -63.9686
0.11363 | -64.0738
0.11393 | -64.1644
0.11424 | -64.1366
0.11455 | -64.0716
0.11486 | -63.9463
0.11517 | -63.7579
0.11548 | -63.5586
0.11579 | -63.307
0.1161 | -63.0993
0.1164 | -62.9004
0.11671 | -62.7493
0.11702 | -62.5782
0.11733 | -62.4021
0.11764 | -62.2897
0.11795 | -62.2212
0.11826 | -62.1586
0.11857 | -62.0975
0.11888 | -62.053
0.11918 | -62.0017
0.11949 | -61.9581
0.1198 | -61.9419
0.12011 | -61.9148
0.12042 | -61.8886
0.12073 | -61.8396
0.12104 | -61.7965
0.12135 | -61.7631
0.12165 | -61.7415
0.12196 | -61.7023
0.12227 | -61.6934
0.12258 | -61.7192
0.12289 | -61.7867
0.1232 | -61.8507
0.12351 | -61.9216
0.12382 | -62.0544
0.12412 | -62.1983
0.12443 | -62.3392
0.12474 | -62.4618
0.12505 | -62.5762
0.12536 | -62.7103
0.12567 | -62.82
0.12598 | -62.8912
0.12629 | -62.9612
0.1266 | -63.0169
0.1269 | -62.9814
0.12721 | -62.9149
0.12752 | -62.8522
0.12783 | -62.8103
0.12814 | -62.7616
0.12845 | -62.7256
0.12876 | -62.6768
0.12907 | -62.6237
0.12937 | -62.5662
0.12968 | -62.5379
0.12999 | -62.5418
0.1303 | -62.5612
0.13061 | -62.5791
0.13092 | -62.6123
0.13123 | -62.6631
0.13154 | -62.7413
0.13184 | -62.835
0.13215 | -62.9472
0.13246 | -63.0924
0.13277 | -63.2502
0.13308 | -63.3938
0.13339 | -63.4506
0.1337 | -63.5406
0.13401 | -63.6391
0.13431 | -63.7161
0.13462 | -63.7575
0.13493 | -63.7787
0.13524 | -63.7507
0.13555 | -63.7194
0.13586 | -63.6894
0.13617 | -63.6626
0.13648 | -63.614
0.13678 | -63.629
0.13709 | -63.6668
0.1374 | -63.7017
0.13771 | -63.7517
0.13802 | -63.825
0.13833 | -63.9227
0.13864 | -64.0296
0.13895 | -64.1313
0.13925 | -64.2717
0.13956 | -64.3923
0.13987 | -64.5023
0.14018 | -64.6009
0.14049 | -64.6794
0.1408 | -64.7624
0.14111 | -64.8073
0.14142 | -64.8282
0.14172 | -64.8362
0.14203 | -64.8319
0.14234 | -64.8152
0.14265 | -64.821
0.14296 | -64.8373
0.14327 | -64.8146
0.14358 | -64.7946
0.14389 | -64.7832
0.14419 | -64.7515
0.1445 | -64.7352
0.14481 | -64.7224
0.14512 | -64.7289
0.14543 | -64.7481
0.14574 | -64.7824
0.14605 | -64.8164
0.14636 | -64.8516
0.14666 | -64.9107
0.14697 | -65.0072
0.14728 | -65.0665
0.14759 | -65.1382
0.1479 | -65.2162
0.14821 | -65.2939
0.14852 | -65.3945
0.14883 | -65.4983
0.14913 | -65.624
0.14944 | -65.7526
0.14975 | -65.8637
0.15006 | -65.9697
0.15037 | -66.0756
0.15068 | -66.1878
0.15099 | -66.3145
0.1513 | -66.4051
0.1516 | -66.4999
0.15191 | -66.6314
0.15222 | -66.7371
0.15253 | -66.8468
0.15284 | -66.9565
0.15315 | -67.0487
0.15346 | -67.1655
0.15377 | -67.2807
0.15407 | -67.4058
0.15438 | -67.5541
0.15469 | -67.7119
0.155 | -67.8237
0.15531 | -67.9164
0.15562 | -68.0081
0.15593 | -68.119
0.15624 | -68.2088
0.15654 | -68.2964
0.15685 | -68.3623
0.15716 | -68.4324
0.15747 | -68.4957
0.15778 | -68.5212
0.15809 | -68.5434
0.1584 | -68.5592
0.15871 | -68.5098
0.15901 | -68.4761
0.15932 | -68.4593
0.15963 | -68.4141
0.15994 | -68.327
0.16025 | -68.3323
0.16056 | -68.3261
0.16087 | -68.3455
0.16118 | -68.3483
0.16149 | -68.3599
0.16179 | -68.3895
0.1621 | -68.4306
0.16241 | -68.4707
0.16272 | -68.5242
0.16303 | -68.5679
0.16334 | -68.5891
0.16365 | -68.6134
0.16396 | -68.662
0.16426 | -68.7038
0.16457 | -68.7447
0.16488 | -68.7839
0.16519 | -68.8137
0.1655 | -68.8546
0.16581 | -68.8811
0.16612 | -68.9275
0.16643 | -69.0003
0.16673 | -69.0577
0.16704 | -69.0881
0.16735 | -69.1142
0.16766 | -69.1277
0.16797 | -69.1347
0.16828 | -69.1399
0.16859 | -69.1671
0.1689 | -69.1852
0.1692 | -69.1696
0.16951 | -69.1608
0.16982 | -69.1516
0.17013 | -69.123
0.17044 | -69.0846
0.17075 | -69.08
0.17106 | -69.0481
0.17137 | -69.0037
0.17167 | -68.9763
0.17198 | -68.9922
0.17229 | -68.9476
0.1726 | -68.9252
0.17291 | -68.8989
0.17322 | -68.8798
0.17353 | -68.8865
0.17384 | -68.9094
0.17414 | -68.9265
0.17445 | -68.9309
0.17476 | -68.9615
0.17507 | -69.0129
0.17538 | -69.0596
0.17569 | -69.1186
0.176 | -69.1678
0.17631 | -69.2071
0.17661 | -69.2478
0.17692 | -69.2749
0.17723 | -69.3124
0.17754 | -69.3434
0.17785 | -69.3558
0.17816 | -69.3588
0.17847 | -69.3987
0.17878 | -69.4251
0.17908 | -69.4273
0.17939 | -69.4476
0.1797 | -69.4841
0.18001 | -69.4906
0.18032 | -69.5028
0.18063 | -69.5256
0.18094 | -69.5446
0.18125 | -69.5507
0.18155 | -69.546
0.18186 | -69.5606
0.18217 | -69.5701
0.18248 | -69.5839
0.18279 | -69.6093
0.1831 | -69.65
0.18341 | -69.6966
0.18372 | -69.7182
0.18403 | -69.7605
0.18433 | -69.8021
0.18464 | -69.8435
0.18495 | -69.8784
0.18526 | -69.9245
0.18557 | -69.9678
0.18588 | -70.0178
0.18619 | -70.0769
0.1865 | -70.1303
0.1868 | -70.1992
0.18711 | -70.2585
0.18742 | -70.3332
0.18773 | -70.4343
0.18804 | -70.515
0.18835 | -70.6023
0.18866 | -70.7022
0.18897 | -70.8137
0.18927 | -70.91
0.18958 | -71.0129
0.18989 | -71.1285
0.1902 | -71.2729
0.19051 | -71.402
0.19082 | -71.5394
0.19113 | -71.687
0.19144 | -71.8105
0.19174 | -71.8696
0.19205 | -71.9391
0.19236 | -72.0147
0.19267 | -72.0824
0.19298 | -72.0946
0.19329 | -72.0829
0.1936 | -72.0687
0.19391 | -71.982
0.19421 | -71.865
0.19452 | -71.7431
0.19483 | -71.6144
0.19514 | -71.5032
0.19545 | -71.3484
0.19576 | -71.1236
0.19607 | -70.9694
0.19638 | -70.7914
0.19668 | -70.6124
0.19699 | -70.4443
0.1973 | -70.2763
0.19761 | -70.1291
0.19792 | -69.9933
0.19823 | -69.8645
0.19854 | -69.7386
0.19885 | -69.6252
0.19915 | -69.5046
0.19946 | -69.3581
0.19977 | -69.2554
0.20008 | -69.183
0.20039 | -69.1117
0.2007 | -69.0584
0.20101 | -69.0161
0.20132 | -68.9753
0.20162 | -68.9361
0.20193 | -68.8963
0.20224 | -68.8649
0.20255 | -68.8351
0.20286 | -68.8134
0.20317 | -68.7903
0.20348 | -68.7739
0.20379 | -68.742
0.20409 | -68.7027
0.2044 | -68.6745
0.20471 | -68.6906
0.20502 | -68.6878
0.20533 | -68.6994
0.20564 | -68.7196
0.20595 | -68.7559
0.20626 | -68.8177
0.20657 | -68.8765
0.20687 | -68.9429
0.20718 | -69.0279
0.20749 | -69.1419
0.2078 | -69.2548
0.20811 | -69.3772
0.20842 | -69.5376
0.20873 | -69.714
0.20904 | -69.8902
0.20935 | -70.0781
0.20965 | -70.2728
0.20996 | -70.4804
0.21027 | -70.7017
0.21058 | -70.9375
0.21089 | -71.192
0.2112 | -71.4392
0.21151 | -71.6631
0.21182 | -71.8884
0.21212 | -72.1303
0.21243 | -72.3641
0.21274 | -72.5823
0.21305 | -72.8039
0.21336 | -73.0433
0.21367 | -73.1552
0.21398 | -73.3597
0.21429 | -73.5841
0.21459 | -73.8427
0.2149 | -73.9744
0.21521 | -74.0376
0.21552 | -74.1132
0.21583 | -74.1536
0.21614 | -74.1734
0.21645 | -74.1673
0.21676 | -74.1522
0.21706 | -74.1117
0.21737 | -74.0708
0.21768 | -74.0137
0.21799 | -73.9244
0.2183 | -73.8523
0.21861 | -73.7248
0.21892 | -73.5768
0.21923 | -73.4655
0.21953 | -73.3306
0.21984 | -73.1844
0.22015 | -73.0322
0.22046 | -72.8782
0.22077 | -72.7231
0.22108 | -72.5691
0.22139 | -72.4211
0.2217 | -72.2828
0.222 | -72.1414
0.22231 | -72.0049
0.22262 | -71.9152
0.22293 | -71.82
0.22324 | -71.7262
0.22355 | -71.6486
0.22386 | -71.5575
0.22417 | -71.4774
0.22447 | -71.3935
0.22478 | -71.3209
0.22509 | -71.2793
0.2254 | -71.2509
0.22571 | -71.2351
0.22602 | -71.226
0.22633 | -71.2262
0.22664 | -71.2353
0.22694 | -71.2601
0.22725 | -71.292
0.22756 | -71.3275
0.22787 | -71.3669
0.22818 | -71.3878
0.22849 | -71.4242
0.2288 | -71.4686
0.22911 | -71.5118
0.22941 | -71.5592
0.22972 | -71.612
0.23003 | -71.6706
0.23034 | -71.7274
0.23065 | -71.779
0.23096 | -71.8303
0.23127 | -71.8755
0.23158 | -71.9357
0.23189 | -71.9722
0.23219 | -72.0265
0.2325 | -72.0861
0.23281 | -72.1333
0.23312 | -72.1917
0.23343 | -72.261
0.23374 | -72.3046
0.23405 | -72.3558
0.23436 | -72.4147
0.23466 | -72.4593
0.23497 | -72.5315
0.23528 | -72.614
0.23559 | -72.6664
0.2359 | -72.7002
0.23621 | -72.732
0.23652 | -72.7677
0.23683 | -72.8004
0.23713 | -72.8232
0.23744 | -72.8448
0.23775 | -72.8669
0.23806 | -72.8816
0.23837 | -72.9044
0.23868 | -72.9287
0.23899 | -72.9411
0.2393 | -72.9532
0.2396 | -72.9711
0.23991 | -72.9896
0.24022 | -73.0185
0.24053 | -73.0596
0.24084 | -73.1117
0.24115 | -73.1563
0.24146 | -73.2085
0.24177 | -73.2441
0.24207 | -73.2481
0.24238 | -73.2586
0.24269 | -73.2739
0.243 | -73.3035
0.24331 | -73.3446
0.24362 | -73.3787
0.24393 | -73.4141
0.24424 | -73.4612
0.24454 | -73.4984
0.24485 | -73.5255
0.24516 | -73.5684
0.24547 | -73.5784
0.24578 | -73.5723
0.24609 | -73.5678
0.2464 | -73.5558
0.24671 | -73.5176
0.24701 | -73.4763
0.24732 | -73.4466
0.24763 | -73.4159
0.24794 | -73.3751
0.24825 | -73.3232
0.24856 | -73.2685
0.24887 | -73.2202
0.24918 | -73.15
0.24948 | -73.1122
0.24979 | -73.0619
0.2501 | -73.0159
0.25041 | -72.9891
0.25072 | -72.9557
0.25103 | -72.9244
0.25134 | -72.9029
0.25165 | -72.8742
0.25195 | -72.8559
0.25226 | -72.8429
0.25257 | -72.83
0.25288 | -72.8267
0.25319 | -72.827
0.2535 | -72.8309
0.25381 | -72.8469
0.25412 | -72.8595
0.25443 | -72.8746
0.25473 | -72.9073
0.25504 | -72.94
0.25535 | -72.9732
0.25566 | -73.008
0.25597 | -73.0355
0.25628 | -73.0661
0.25659 | -73.1081
0.25689 | -73.1382
0.2572 | -73.1851
0.25751 | -73.2339
0.25782 | -73.2777
0.25813 | -73.3197
0.25844 | -73.3727
0.25875 | -73.4315
0.25906 | -73.482
0.25937 | -73.5332
0.25967 | -73.6001
0.25998 | -73.6667
0.26029 | -73.7278
0.2606 | -73.7857
0.26091 | -73.8331
0.26122 | -73.8755
0.26153 | -73.9092
0.26183 | -73.9456
0.26214 | -73.9819
0.26245 | -74.0235
0.26276 | -74.049
0.26307 | -74.0829
0.26338 | -74.1259
0.26369 | -74.1731
0.264 | -74.221
0.26431 | -74.2707
0.26461 | -74.3179
0.26492 | -74.3854
0.26523 | -74.4485
0.26554 | -74.4989
0.26585 | -74.5537
0.26616 | -74.6141
0.26647 | -74.688
0.26677 | -74.7549
0.26708 | -74.8233
0.26739 | -74.8847
0.2677 | -74.9503
0.26801 | -75.0421
0.26832 | -75.1094
0.26863 | -75.1735
0.26894 | -75.2417
0.26925 | -75.2968
0.26955 | -75.3571
0.26986 | -75.4136
0.27017 | -75.4764
0.27048 | -75.5471
0.27079 | -75.6107
0.2711 | -75.6735
0.27141 | -75.7361
0.27172 | -75.8007
0.27202 | -75.8705
0.27233 | -75.9363
0.27264 | -75.9994
0.27295 | -76.0558
0.27326 | -76.1215
0.27357 | -76.1842
0.27388 | -76.1916
0.27419 | -76.2276
0.27449 | -76.2724
0.2748 | -76.337
0.27511 | -76.3851
0.27542 | -76.4101
0.27573 | -76.4476
0.27604 | -76.4052
0.27635 | -76.4024
0.27666 | -76.3907
0.27697 | -76.3778
0.27727 | -76.3643
0.27758 | -76.3204
0.27789 | -76.2817
0.2782 | -76.2346
0.27851 | -76.1868
0.27882 | -76.1135
0.27913 | -76.0258
0.27943 | -75.9555
0.27974 | -75.8923
0.28005 | -75.8193
0.28036 | -75.7557
0.28067 | -75.6862
0.28098 | -75.6047
0.28129 | -75.5366
0.2816 | -75.4637
0.28191 | -75.3791
0.28221 | -75.3136
0.28252 | -75.26
0.28283 | -75.2005
0.28314 | -75.1462
0.28345 | -75.0913
0.28376 | -75.0392
0.28407 | -75.0003
0.28437 | -74.9675
0.28468 | -74.9357
0.28499 | -74.9409
0.2853 | -74.9149
0.28561 | -74.9003
0.28592 | -74.8933
0.28623 | -74.8886
0.28654 | -74.8917
0.28685 | -74.9034
0.28715 | -74.9185
0.28746 | -74.9382
0.28777 | -74.9679
0.28808 | -74.9902
0.28839 | -75.02
0.2887 | -75.0575
0.28901 | -75.0966
0.28931 | -75.1426
0.28962 | -75.1836
0.28993 | -75.2379
0.29024 | -75.293
0.29055 | -75.353
0.29086 | -75.4114
0.29117 | -75.4765
0.29148 | -75.5469
0.29179 | -75.612
0.29209 | -75.6761
0.2924 | -75.7622
0.29271 | -75.8343
0.29302 | -75.9114
0.29333 | -75.9828
0.29364 | -76.0487
0.29395 | -76.0905
0.29426 | -76.1699
0.29456 | -76.2163
0.29487 | -76.2643
0.29518 | -76.3012
0.29549 | -76.3271
0.2958 | -76.3585
0.29611 | -76.3986
0.29642 | -76.4415
0.29673 | -76.4556
0.29703 | -76.4639
0.29734 | -76.4751
0.29765 | -76.4809
0.29796 | -76.4761
0.29827 | -76.4516
0.29858 | -76.4318
0.29889 | -76.4085
0.2992 | -76.3795
0.29951 | -76.3775
0.29981 | -76.3766
0.30012 | -76.3553
0.30043 | -76.326
0.30074 | -76.2862
0.30105 | -76.2462
0.30136 | -76.2204
0.30167 | -76.1983
0.30197 | -76.1743
0.30228 | -76.1606
0.30259 | -76.1512
0.3029 | -76.1469
0.30321 | -76.1532
0.30352 | -76.1355
0.30383 | -76.1211
0.30414 | -76.1228
0.30445 | -76.1021
0.30475 | -76.0686
0.30506 | -76.056
0.30537 | -76.0557
0.30568 | -76.0561
0.30599 | -76.0493
0.3063 | -76.0421
0.30661 | -76.0314
0.30691 | -76.0195
0.30722 | -75.9989
0.30753 | -75.9675
0.30784 | -75.931
0.30815 | -75.9059
0.30846 | -75.8895
0.30877 | -75.8632
0.30908 | -75.8131
0.30939 | -75.7883
0.30969 | -75.7512
0.31 | -75.7125
0.31031 | -75.6747
0.31062 | -75.6349
0.31093 | -75.5768
0.31124 | -75.5112
0.31155 | -75.3839
0.31185 | -75.3155
0.31216 | -75.2364
0.31247 | -75.1431
0.31278 | -75.037
0.31309 | -74.9416
0.3134 | -74.8605
0.31371 | -74.7654
0.31402 | -74.6688
0.31433 | -74.5793
0.31463 | -74.4937
0.31494 | -74.412
0.31525 | -74.324
0.31556 | -74.2423
0.31587 | -74.1665
0.31618 | -74.0979
0.31649 | -74.0245
0.3168 | -73.9528
0.3171 | -73.8893
0.31741 | -73.8359
0.31772 | -73.7831
0.31803 | -73.7368
0.31834 | -73.6965
0.31865 | -73.666
0.31896 | -73.6415
0.31927 | -73.626
0.31957 | -73.6115
0.31988 | -73.608
0.32019 | -73.6201
0.3205 | -73.6816
0.32081 | -73.7013
0.32112 | -73.7276
0.32143 | -73.7574
0.32174 | -73.7948
0.32205 | -73.8408
0.32235 | -73.8964
0.32266 | -73.9573
0.32297 | -74.0252
0.32328 | -74.1016
0.32359 | -74.1799
0.3239 | -74.2671
0.32421 | -74.3557
0.32451 | -74.4549
0.32482 | -74.5642
0.32513 | -74.6819
0.32544 | -74.7977
0.32575 | -74.9042
0.32606 | -75.0255
0.32637 | -75.1573
0.32668 | -75.2852
0.32699 | -75.4275
0.32729 | -75.5766
0.3276 | -75.7053
0.32791 | -75.8346
0.32822 | -75.9594
0.32853 | -76.094
0.32884 | -76.2349
0.32915 | -76.3632
0.32945 | -76.4292
0.32976 | -76.5451
0.33007 | -76.6661
0.33038 | -76.7931
0.33069 | -76.908
0.331 | -76.99
0.33131 | -77.0523
0.33162 | -77.1245
0.33193 | -77.2161
0.33223 | -77.269
0.33254 | -77.3265
0.33285 | -77.3639
0.33316 | -77.3907
0.33347 | -77.4328
0.33378 | -77.4743
0.33409 | -77.4979
0.33439 | -77.5401
0.3347 | -77.5964
0.33501 | -77.5793
0.33532 | -77.5783
0.33563 | -77.6144
0.33594 | -77.653
0.33625 | -77.6644
0.33656 | -77.656
0.33687 | -77.661
0.33717 | -77.6446
0.33748 | -77.6247
0.33779 | -77.6087
0.3381 | -77.589
0.33841 | -77.5736
0.33872 | -77.5539
0.33903 | -77.5334
0.33934 | -77.5028
0.33964 | -77.4543
0.33995 | -77.4241
0.34026 | -77.394
0.34057 | -77.3533
0.34088 | -77.3164
0.34119 | -77.28
0.3415 | -77.2438
0.34181 | -77.2059
0.34212 | -77.1585
0.34242 | -77.1124
0.34273 | -77.0599
0.34304 | -77.0294
0.34335 | -76.9925
0.34366 | -76.9462
0.34397 | -76.894
0.34428 | -76.8153
0.34458 | -76.7228
0.34489 | -76.658
0.3452 | -76.6083
0.34551 | -76.545
0.34582 | -76.476
0.34613 | -76.4061
0.34644 | -76.3401
0.34675 | -76.2895
0.34706 | -76.2331
0.34736 | -76.173
0.34767 | -76.1197
0.34798 | -76.0682
0.34829 | -76.0145
0.3486 | -75.9578
0.34891 | -75.8973
0.34922 | -75.831
0.34952 | -75.7756
0.34983 | -75.7325
0.35014 | -75.6814
0.35045 | -75.6208
0.35076 | -75.5595
0.35107 | -75.507
0.35138 | -75.4612
0.35169 | -75.4224
0.352 | -75.3814
0.3523 | -75.3467
0.35261 | -75.3128
0.35292 | -75.281
0.35323 | -75.2548
0.35354 | -75.2279
0.35385 | -75.2011
0.35416 | -75.1796
0.35446 | -75.1598
0.35477 | -75.1404
0.35508 | -75.128
0.35539 | -75.1257
0.3557 | -75.1223
0.35601 | -75.119
0.35632 | -75.1513
0.35663 | -75.1584
0.35694 | -75.1687
0.35724 | -75.1807
0.35755 | -75.2005
0.35786 | -75.2311
0.35817 | -75.2548
0.35848 | -75.275
0.35879 | -75.2949
0.3591 | -75.319
0.35941 | -75.3481
0.35971 | -75.3981
0.36002 | -75.4439
0.36033 | -75.471
0.36064 | -75.4976
0.36095 | -75.5288
0.36126 | -75.561
0.36157 | -75.5911
0.36188 | -75.6319
0.36218 | -75.6726
0.36249 | -75.704
0.3628 | -75.7357
0.36311 | -75.767
0.36342 | -75.7988
0.36373 | -75.8341
0.36404 | -75.8607
0.36435 | -75.8875
0.36466 | -75.9278
0.36496 | -75.9217
0.36527 | -75.941
0.36558 | -75.956
0.36589 | -75.969
0.3662 | -75.9812
0.36651 | -75.9868
0.36682 | -75.9753
0.36712 | -75.9618
0.36743 | -75.9612
0.36774 | -75.9608
0.36805 | -75.9606
0.36836 | -75.9526
0.36867 | -75.9441
0.36898 | -75.9344
0.36929 | -75.9159
0.3696 | -75.9189
0.3699 | -75.9304
0.37021 | -75.9234
0.37052 | -75.9081
0.37083 | -75.8909
0.37114 | -75.8667
0.37145 | -75.8532
0.37176 | -75.845
0.37206 | -75.8318
0.37237 | -75.8222
0.37268 | -75.8163
0.37299 | -75.8117
0.3733 | -75.8027
0.37361 | -75.7925
0.37392 | -75.8089
0.37423 | -75.8152
0.37454 | -75.8186
0.37484 | -75.8286
0.37515 | -75.8492
0.37546 | -75.8668
0.37577 | -75.8843
0.37608 | -75.9033
0.37639 | -75.9275
0.3767 | -75.9552
0.377 | -75.9776
0.37731 | -76.0023
0.37762 | -76.0281
0.37793 | -76.0564
0.37824 | -76.0858
0.37855 | -76.1188
0.37886 | -76.1575
0.37917 | -76.1995
0.37948 | -76.246
0.37978 | -76.2918
0.38009 | -76.3377
0.3804 | -76.3905
0.38071 | -76.4479
0.38102 | -76.5062
0.38133 | -76.5652
0.38164 | -76.6296
0.38195 | -76.6973
0.38225 | -76.7577
0.38256 | -76.8244
0.38287 | -76.902
0.38318 | -76.9665
0.38349 | -77.038
0.3838 | -77.1163
0.38411 | -77.1944
0.38442 | -77.2746
0.38472 | -77.3574
0.38503 | -77.4521
0.38534 | -77.5487
0.38565 | -77.637
0.38596 | -77.7308
0.38627 | -77.8163
0.38658 | -77.9103
0.38689 | -78.016
0.3872 | -78.1229
0.3875 | -78.2093
0.38781 | -78.2818
0.38812 | -78.389
0.38843 | -78.4753
0.38874 | -78.5554
0.38905 | -78.641
0.38936 | -78.7206
0.38966 | -78.7816
0.38997 | -78.8448
0.39028 | -78.9078
0.39059 | -78.9775
0.3909 | -79.0409
0.39121 | -79.0903
0.39152 | -79.1458
0.39183 | -79.1452
0.39214 | -79.1852
0.39244 | -79.2185
0.39275 | -79.2416
0.39306 | -79.2673
0.39337 | -79.2684
0.39368 | -79.267
0.39399 | -79.2801
0.3943 | -79.29
0.3946 | -79.2805
0.39491 | -79.2579
0.39522 | -79.2311
0.39553 | -79.1889
0.39584 | -79.1214
0.39615 | -79.0635
0.39646 | -79.0362
0.39677 | -78.9972
0.39708 | -78.938
0.39738 | -78.8679
0.39769 | -78.8085
0.398 | -78.7626
0.39831 | -78.7263
0.39862 | -78.6698
0.39893 | -78.605
0.39924 | -78.5385
0.39954 | -78.4759
0.39985 | -78.4141
0.40016 | -78.3466
0.40047 | -78.2754
0.40078 | -78.2102
0.40109 | -78.1668
0.4014 | -78.1166
0.40171 | -78.0587
0.40202 | -78.0082
0.40232 | -77.9626
0.40263 | -77.9156
0.40294 | -77.8695
0.40325 | -77.8283
0.40356 | -77.7917
0.40387 | -77.7573
0.40418 | -77.7301
0.40449 | -77.7024
0.4048 | -77.6691
0.4051 | -77.6406
0.40541 | -77.617
0.40572 | -77.5896
0.40603 | -77.5736
0.40634 | -77.5565
0.40665 | -77.5251
0.40696 | -77.5117
0.40726 | -77.5031
0.40757 | -77.4971
0.40788 | -77.4889
0.40819 | -77.4863
0.4085 | -77.4796
0.40881 | -77.4761
0.40912 | -77.4809
0.40943 | -77.4858
0.40974 | -77.4894
0.41004 | -77.4919
0.41035 | -77.5008
0.41066 | -77.5099
0.41097 | -77.5227
0.41128 | -77.5404
0.41159 | -77.5602
0.4119 | -77.5626
0.4122 | -77.5504
0.41251 | -77.558
0.41282 | -77.5678
0.41313 | -77.5808
0.41344 | -77.5941
0.41375 | -77.6104
0.41406 | -77.6275
0.41437 | -77.6338
0.41468 | -77.6391
0.41498 | -77.6496
0.41529 | -77.659
0.4156 | -77.6696
0.41591 | -77.6916
0.41622 | -77.7209
0.41653 | -77.7266
0.41684 | -77.7372
0.41714 | -77.7451
0.41745 | -77.754
0.41776 | -77.7672
0.41807 | -77.7853
0.41838 | -77.794
0.41869 | -77.8005
0.419 | -77.8061
0.41931 | -77.8102
0.41962 | -77.8227
0.41992 | -77.8356
0.42023 | -77.8467
0.42054 | -77.8531
0.42085 | -77.8653
0.42116 | -77.8775
0.42147 | -77.8871
0.42178 | -77.8957
0.42209 | -77.9078
0.42239 | -77.9281
0.4227 | -77.9503
0.42301 | -77.9777
0.42332 | -77.9831
0.42363 | -77.9912
0.42394 | -78.0085
0.42425 | -78.0235
0.42456 | -78.041
0.42486 | -78.0815
0.42517 | -78.1084
0.42548 | -78.1165
0.42579 | -78.13
0.4261 | -78.151
0.42641 | -78.1753
0.42672 | -78.1941
0.42703 | -78.2104
0.42734 | -78.2223
0.42764 | -78.2354
0.42795 | -78.2537
0.42826 | -78.2707
0.42857 | -78.2875
0.42888 | -78.297
0.42919 | -78.3087
0.4295 | -78.33
0.4298 | -78.358
0.43011 | -78.3843
0.43042 | -78.3937
0.43073 | -78.4014
0.43104 | -78.4101
0.43135 | -78.4187
0.43166 | -78.4232
0.43197 | -78.4157
0.43228 | -78.4108
0.43258 | -78.4044
0.43289 | -78.395
0.4332 | -78.388
0.43351 | -78.3826
0.43382 | -78.3757
0.43413 | -78.3743
0.43444 | -78.3735
0.43474 | -78.3496
0.43505 | -78.3369
0.43536 | -78.3364
0.43567 | -78.3265
0.43598 | -78.319
0.43629 | -78.3146
0.4366 | -78.3195
0.43691 | -78.3001
0.43722 | -78.2768
0.43752 | -78.265
0.43783 | -78.2578
0.43814 | -78.2345
0.43845 | -78.2252
0.43876 | -78.2171
0.43907 | -78.2023
0.43938 | -78.1924
0.43968 | -78.1906
0.43999 | -78.1912
0.4403 | -78.1864
0.44061 | -78.1791
0.44092 | -78.1742
0.44123 | -78.1697
0.44154 | -78.1664
0.44185 | -78.1614
0.44216 | -78.1709
0.44246 | -78.1921
0.44277 | -78.1907
0.44308 | -78.1829
0.44339 | -78.1724
0.4437 | -78.1631
0.44401 | -78.1583
0.44432 | -78.1539
0.44463 | -78.1574
0.44493 | -78.1631
0.44524 | -78.1711
0.44555 | -78.1722
0.44586 | -78.1733
0.44617 | -78.1679
0.44648 | -78.1584
0.44679 | -78.1551
0.4471 | -78.1591
0.4474 | -78.165
0.44771 | -78.1713
0.44802 | -78.1749
0.44833 | -78.1802
0.44864 | -78.1839
0.44895 | -78.1908
0.44926 | -78.1985
0.44957 | -78.1984
0.44988 | -78.2005
0.45018 | -78.2062
0.45049 | -78.2133
0.4508 | -78.2215
0.45111 | -78.2301
0.45142 | -78.2428
0.45173 | -78.2523
0.45204 | -78.2628
0.45234 | -78.2749
0.45265 | -78.2871
0.45296 | -78.2983
0.45327 | -78.3127
0.45358 | -78.3271
0.45389 | -78.3376
0.4542 | -78.3537
0.45451 | -78.3729
0.45482 | -78.383
0.45512 | -78.3877
0.45543 | -78.4017
0.45574 | -78.4234
0.45605 | -78.4343
0.45636 | -78.4502
0.45667 | -78.4741
0.45698 | -78.4893
0.45728 | -78.5026
0.45759 | -78.5146
0.4579 | -78.519
0.45821 | -78.5308
0.45852 | -78.5517
0.45883 | -78.5732
0.45914 | -78.5933
0.45945 | -78.6079
0.45976 | -78.6259
0.46006 | -78.6434
0.46037 | -78.6661
0.46068 | -78.6888
0.46099 | -78.7085
0.4613 | -78.7252
0.46161 | -78.7391
0.46192 | -78.7551
0.46222 | -78.7718
0.46253 | -78.7866
0.46284 | -78.8081
0.46315 | -78.829
0.46346 | -78.8445
0.46377 | -78.872
0.46408 | -78.8963
0.46439 | -78.9204
0.4647 | -78.9458
0.465 | -78.9664
0.46531 | -78.9869
0.46562 | -79.0079
0.46593 | -79.0265
0.46624 | -79.05
0.46655 | -79.0911
0.46686 | -79.111
0.46717 | -79.1235
0.46747 | -79.132
0.46778 | -79.1332
0.46809 | -79.138
0.4684 | -79.1503
0.46871 | -79.1625
0.46902 | -79.1824
0.46933 | -79.2069
0.46964 | -79.2055
0.46995 | -79.213
0.47025 | -79.232
0.47056 | -79.2444
0.47087 | -79.2458
0.47118 | -79.2534
0.47149 | -79.2586
0.4718 | -79.2662
0.47211 | -79.2649
0.47241 | -79.2673
0.47272 | -79.2684
0.47303 | -79.2715
0.47334 | -79.2708
0.47365 | -79.2717
0.47396 | -79.2766
0.47427 | -79.2787
0.47458 | -79.2829
0.47489 | -79.2899
0.47519 | -79.2961
0.4755 | -79.2945
0.47581 | -79.2931
0.47612 | -79.2946
0.47643 | -79.2978
0.47674 | -79.2926
0.47705 | -79.2761
0.47735 | -79.2659
0.47766 | -79.254
0.47797 | -79.2433
0.47828 | -79.2336
0.47859 | -79.2243
0.4789 | -79.2168
0.47921 | -79.2161
0.47952 | -79.2108
0.47983 | -79.1983
0.48013 | -79.1866
0.48044 | -79.1744
0.48075 | -79.1654
0.48106 | -79.1482
0.48137 | -79.1362
0.48168 | -79.1199
0.48199 | -79.0981
0.48229 | -79.0886
0.4826 | -79.0868
0.48291 | -79.0767
0.48322 | -79.0567
0.48353 | -79.04
0.48384 | -79.0267
0.48415 | -79.0136
0.48446 | -78.9964
0.48477 | -78.9776
0.48507 | -78.9663
0.48538 | -78.9591
0.48569 | -78.9532
0.486 | -78.9425
0.48631 | -78.9329
0.48662 | -78.9279
0.48693 | -78.9165
0.48724 | -78.9133
0.48754 | -78.9086
0.48785 | -78.9126
0.48816 | -78.9134
0.48847 | -78.909
0.48878 | -78.9055
0.48909 | -78.9018
0.4894 | -78.8968
0.48971 | -78.9014
0.49001 | -78.9042
0.49032 | -78.9056
0.49063 | -78.9042
0.49094 | -78.906
0.49125 | -78.9141
0.49156 | -78.9229
0.49187 | -78.9266
0.49218 | -78.9307
0.49249 | -78.9316
0.49279 | -78.9293
0.4931 | -78.9342
0.49341 | -78.9424
0.49372 | -78.953
0.49403 | -78.9603
0.49434 | -78.9647
0.49465 | -78.9713
0.49495 | -78.9855
0.49526 | -78.9991
0.49557 | -79.0083
0.49588 | -79.0207
0.49619 | -79.0292
0.4965 | -79.0366
0.49681 | -79.0493
0.49712 | -79.0588
0.49743 | -79.0722
0.49773 | -79.0887
0.49804 | -79.107
0.49835 | -79.1213
0.49866 | -79.1414
0.49897 | -79.1555
0.49928 | -79.1659
0.49959 | -79.1787
0.49989 | -79.1911
0.5002 | -79.2062
0.50051 | -79.2245
0.50082 | -79.2421
0.50113 | -79.2623
0.50144 | -79.2799
0.50175 | -79.3036
0.50206 | -79.333
0.50236 | -79.3586
0.50267 | -79.378
0.50298 | -79.4015
0.50329 | -79.4251
0.5036 | -79.4445
0.50391 | -79.4761
0.50422 | -79.498
0.50453 | -79.5167
0.50484 | -79.5365
0.50514 | -79.5636
0.50545 | -79.5846
0.50576 | -79.6071
0.50607 | -79.6242
0.50638 | -79.6459
0.50669 | -79.6736
0.507 | -79.6999
0.50731 | -79.725
0.50761 | -79.7562
0.50792 | -79.7828
0.50823 | -79.8136
0.50854 | -79.8439
0.50885 | -79.8642
0.50916 | -79.8864
0.50947 | -79.9143
0.50978 | -79.9391
0.51008 | -79.9681
0.51039 | -80.0106
0.5107 | -80.0517
0.51101 | -80.0902
0.51132 | -80.1245
0.51163 | -80.1576
0.51194 | -80.1939
0.51225 | -80.2199
0.51255 | -80.2541
0.51286 | -80.2851
0.51317 | -80.315
0.51348 | -80.3466
0.51379 | -80.3813
0.5141 | -80.4152
0.51441 | -80.4439
0.51472 | -80.4766
0.51502 | -80.5065
0.51533 | -80.54
0.51564 | -80.5694
0.51595 | -80.5949
0.51626 | -80.6173
0.51657 | -80.6308
0.51688 | -80.6575
0.51719 | -80.6928
0.5175 | -80.7302
0.5178 | -80.7544
0.51811 | -80.7765
0.51842 | -80.8015
0.51873 | -80.8344
0.51904 | -80.8588
0.51935 | -80.8783
0.51966 | -80.8966
0.51997 | -80.9198
0.52027 | -80.9451
0.52058 | -80.962
0.52089 | -80.974
0.5212 | -80.9834
0.52151 | -80.9993
0.52182 | -81.0135
0.52213 | -81.0205
0.52243 | -81.0402
0.52274 | -81.0485
0.52305 | -81.0621
0.52336 | -81.0722
0.52367 | -81.0796
0.52398 | -81.0869
0.52429 | -81.092
0.5246 | -81.0943
0.5249 | -81.106
0.52521 | -81.1178
0.52552 | -81.1141
0.52583 | -81.1124
0.52614 | -81.1111
0.52645 | -81.1111
0.52676 | -81.1105
0.52707 | -81.112
0.52738 | -81.1095
0.52768 | -81.1175
0.52799 | -81.1241
0.5283 | -81.1245
0.52861 | -81.1341
0.52892 | -81.1425
0.52923 | -81.1356
0.52954 | -81.1247
0.52985 | -81.1247
0.53015 | -81.1323
0.53046 | -81.1367
0.53077 | -81.1251
0.53108 | -81.0995
0.53139 | -81.0841
0.5317 | -81.0861
0.53201 | -81.0829
0.53232 | -81.0856
0.53262 | -81.0806
0.53293 | -81.0677
0.53324 | -81.0666
0.53355 | -81.0639
0.53386 | -81.0566
0.53417 | -81.0447
0.53448 | -81.0309
0.53479 | -81.0223
0.53509 | -81.0099
0.5354 | -80.9938
0.53571 | -80.9769
0.53602 | -80.9715
0.53633 | -80.9564
0.53664 | -80.9426
0.53695 | -80.9298
0.53726 | -80.91
0.53756 | -80.8853
0.53787 | -80.8621
0.53818 | -80.8461
0.53849 | -80.8284
0.5388 | -80.8101
0.53911 | -80.7865
0.53942 | -80.7617
0.53973 | -80.7397
0.54004 | -80.7154
0.54034 | -80.6886
0.54065 | -80.6562
0.54096 | -80.6267
0.54127 | -80.5859
0.54158 | -80.5504
0.54189 | -80.5223
0.5422 | -80.4972
0.54251 | -80.4723
0.54281 | -80.4486
0.54312 | -80.425
0.54343 | -80.404
0.54374 | -80.3797
0.54405 | -80.3578
0.54436 | -80.3361
0.54467 | -80.3127
0.54497 | -80.293
0.54528 | -80.2678
0.54559 | -80.2445
0.5459 | -80.2247
0.54621 | -80.2064
0.54652 | -80.1893
0.54683 | -80.1728
0.54714 | -80.1558
0.54744 | -80.1421
0.54775 | -80.1281
0.54806 | -80.1149
0.54837 | -80.0974
0.54868 | -80.085
0.54899 | -80.078
0.5493 | -80.0771
0.54961 | -80.0795
0.54992 | -80.0813
0.55022 | -80.0813
0.55053 | -80.0791
0.55084 | -80.0828
0.55115 | -80.0886
0.55146 | -80.0915
0.55177 | -80.1002
0.55208 | -80.131
0.55239 | -80.1413
0.55269 | -80.1512
0.553 | -80.1616
0.55331 | -80.1769
0.55362 | -80.2024
0.55393 | -80.2291
0.55424 | -80.2534
0.55455 | -80.2766
0.55486 | -80.2976
0.55516 | -80.3204
0.55547 | -80.3471
0.55578 | -80.3738
0.55609 | -80.3982
0.5564 | -80.4301
0.55671 | -80.4623
0.55702 | -80.4934
0.55733 | -80.5254
0.55763 | -80.5513
0.55794 | -80.5777
0.55825 | -80.6056
0.55856 | -80.633
0.55887 | -80.6682
0.55918 | -80.7076
0.55949 | -80.7411
0.5598 | -80.7718
0.5601 | -80.8093
0.56041 | -80.8432
0.56072 | -80.8688
0.56103 | -80.9021
0.56134 | -80.938
0.56165 | -80.9728
0.56196 | -81.0093
0.56227 | -81.0461
0.56258 | -81.071
0.56288 | -81.0938
0.56319 | -81.1159
0.5635 | -81.1468
0.56381 | -81.1803
0.56412 | -81.2119
0.56443 | -81.232
0.56474 | -81.2542
0.56505 | -81.2737
0.56535 | -81.2883
0.56566 | -81.3019
0.56597 | -81.3196
0.56628 | -81.3363
0.56659 | -81.3457
0.5669 | -81.3501
0.56721 | -81.3604
0.56751 | -81.3766
0.56782 | -81.3951
0.56813 | -81.4046
0.56844 | -81.404
0.56875 | -81.398
0.56906 | -81.3967
0.56937 | -81.4092
0.56968 | -81.4278
0.56998 | -81.4049
0.57029 | -81.4007
0.5706 | -81.3859
0.57091 | -81.3735
0.57122 | -81.3592
0.57153 | -81.3479
0.57184 | -81.3335
0.57215 | -81.3157
0.57246 | -81.2848
0.57276 | -81.2564
0.57307 | -81.238
0.57338 | -81.2176
0.57369 | -81.1809
0.574 | -81.1443
0.57431 | -81.1181
0.57462 | -81.0949
0.57493 | -81.0689
0.57523 | -81.0347
0.57554 | -81.008
0.57585 | -80.9834
0.57616 | -80.9551
0.57647 | -80.9295
0.57678 | -80.9078
0.57709 | -80.8852
0.5774 | -80.8613
0.5777 | -80.8343
0.57801 | -80.8088
0.57832 | -80.7794
0.57863 | -80.7499
0.57894 | -80.7201
0.57925 | -80.6865
0.57956 | -80.6618
0.57987 | -80.6418
0.58017 | -80.6276
0.58048 | -80.61
0.58079 | -80.6043
0.5811 | -80.5901
0.58141 | -80.5692
0.58172 | -80.5506
0.58203 | -80.5297
0.58234 | -80.5112
0.58264 | -80.4953
0.58295 | -80.4755
0.58326 | -80.4587
0.58357 | -80.4476
0.58388 | -80.438
0.58419 | -80.4297
0.5845 | -80.4225
0.58481 | -80.4189
0.58512 | -80.4178
0.58542 | -80.4173
0.58573 | -80.4117
0.58604 | -80.4164
0.58635 | -80.4262
0.58666 | -80.4282
0.58697 | -80.4339
0.58728 | -80.4385
0.58759 | -80.4384
0.58789 | -80.4471
0.5882 | -80.4563
0.58851 | -80.4686
0.58882 | -80.4774
0.58913 | -80.4868
0.58944 | -80.497
0.58975 | -80.5094
0.59005 | -80.5225
0.59036 | -80.5356
0.59067 | -80.5556
0.59098 | -80.5666
0.59129 | -80.5789
0.5916 | -80.5995
0.59191 | -80.6165
0.59222 | -80.632
0.59252 | -80.6527
0.59283 | -80.6725
0.59314 | -80.6913
0.59345 | -80.7165
0.59376 | -80.7427
0.59407 | -80.767
0.59438 | -80.7904
0.59469 | -80.8137
0.595 | -80.8323
0.5953 | -80.8553
0.59561 | -80.8807
0.59592 | -80.9066
0.59623 | -80.9339
0.59654 | -80.964
0.59685 | -80.9915
0.59716 | -81.0166
0.59747 | -81.0432
0.59777 | -81.0731
0.59808 | -81.0983
0.59839 | -81.1318
0.5987 | -81.1699
0.59901 | -81.2056
0.59932 | -81.2415
0.59963 | -81.2708
0.59994 | -81.2993
0.60025 | -81.3319
0.60055 | -81.3644
0.60086 | -81.39
0.60117 | -81.4159
0.60148 | -81.4544
0.60179 | -81.4831
0.6021 | -81.5093
0.60241 | -81.5383
0.60272 | -81.5683
0.60302 | -81.5962
0.60333 | -81.6245
0.60364 | -81.6626
0.60395 | -81.7032
0.60426 | -81.7362
0.60457 | -81.7519
0.60488 | -81.7646
0.60518 | -81.7823
0.60549 | -81.8012
0.6058 | -81.8346
0.60611 | -81.8524
0.60642 | -81.8646
0.60673 | -81.8899
0.60704 | -81.9129
0.60735 | -81.9267
0.60765 | -81.9416
0.60796 | -81.9618
0.60827 | -81.9819
0.60858 | -81.9779
0.60889 | -81.9808
0.6092 | -81.9903
0.60951 | -82.0055
0.60982 | -82.0214
0.61013 | -82.033
0.61043 | -82.0391
0.61074 | -82.0466
0.61105 | -82.0545
0.61136 | -82.061
0.61167 | -82.0691
0.61198 | -82.0782
0.61229 | -82.0841
0.6126 | -82.0875
0.6129 | -82.0935
0.61321 | -82.0932
0.61352 | -82.0951
0.61383 | -82.1002
0.61414 | -82.1117
0.61445 | -82.1057
0.61476 | -82.1112
0.61506 | -82.1012
0.61537 | -82.0958
0.61568 | -82.1047
0.61599 | -82.1138
0.6163 | -82.112
0.61661 | -82.1055
0.61692 | -82.1033
0.61723 | -82.1037
0.61754 | -82.0986
0.61784 | -82.09
0.61815 | -82.0761
0.61846 | -82.0583
0.61877 | -82.0431
0.61908 | -82.0286
0.61939 | -82.0154
0.6197 | -82.0074
0.62001 | -82.0008
0.62031 | -81.9967
0.62062 | -81.9946
0.62093 | -81.9914
0.62124 | -81.9944
0.62155 | -81.9947
0.62186 | -81.9825
0.62217 | -81.971
0.62248 | -81.961
0.62279 | -81.948
0.62309 | -81.9402
0.6234 | -81.9344
0.62371 | -81.9282
0.62402 | -81.9211
0.62433 | -81.9141
0.62464 | -81.9141
0.62495 | -81.9155
0.62526 | -81.9107
0.62556 | -81.9142
0.62587 | -81.9123
0.62618 | -81.9085
0.62649 | -81.9044
0.6268 | -81.8988
0.62711 | -81.895
0.62742 | -81.8902
0.62772 | -81.8814
0.62803 | -81.8757
0.62834 | -81.8814
0.62865 | -81.8847
0.62896 | -81.8818
0.62927 | -81.8788
0.62958 | -81.8807
0.62989 | -81.8827
0.63019 | -81.8943
0.6305 | -81.9039
0.63081 | -81.9104
0.63112 | -81.9151
0.63143 | -81.9226
0.63174 | -81.9312
0.63205 | -81.9382
0.63236 | -81.9497
0.63267 | -81.9569
0.63297 | -81.9604
0.63328 | -81.9628
0.63359 | -81.9712
0.6339 | -81.9774
0.63421 | -81.9952
0.63452 | -82.0117
0.63483 | -82.0204
0.63514 | -82.0307
0.63544 | -82.0414
0.63575 | -82.0474
0.63606 | -82.0533
0.63637 | -82.0689
0.63668 | -82.0791
0.63699 | -82.0899
0.6373 | -82.1005
0.6376 | -82.111
0.63791 | -82.1215
0.63822 | -82.1358
0.63853 | -82.1416
0.63884 | -82.1461
0.63915 | -82.1568
0.63946 | -82.1663
0.63977 | -82.1743
0.64008 | -82.1846
0.64038 | -82.1931
0.64069 | -82.2009
0.641 | -82.2093
0.64131 | -82.2171
0.64162 | -82.2255
0.64193 | -82.2381
0.64224 | -82.2516
0.64255 | -82.26
0.64285 | -82.2676
0.64316 | -82.2739
0.64347 | -82.2811
0.64378 | -82.2898
0.64409 | -82.303
0.6444 | -82.314
0.64471 | -82.3249
0.64502 | -82.3331
0.64533 | -82.3433
0.64563 | -82.3534
0.64594 | -82.3636
0.64625 | -82.3618
0.64656 | -82.3701
0.64687 | -82.3816
0.64718 | -82.3853
0.64749 | -82.3899
0.6478 | -82.3958
0.6481 | -82.4047
0.64841 | -82.4222
0.64872 | -82.4383
0.64903 | -82.4573
0.64934 | -82.4686
0.64965 | -82.4745
0.64996 | -82.4852
0.65026 | -82.4946
0.65057 | -82.5134
0.65088 | -82.5334
0.65119 | -82.5446
0.6515 | -82.5517
0.65181 | -82.5592
0.65212 | -82.5661
0.65243 | -82.5744
0.65273 | -82.5815
0.65304 | -82.6062
0.65335 | -82.6212
0.65366 | -82.6271
0.65397 | -82.631
0.65428 | -82.6345
0.65459 | -82.6376
0.6549 | -82.6461
0.65521 | -82.6532
0.65551 | -82.6603
0.65582 | -82.6686
0.65613 | -82.6749
0.65644 | -82.6851
0.65675 | -82.6927
0.65706 | -82.7008
0.65737 | -82.7079
0.65768 | -82.7183
0.65798 | -82.7302
0.65829 | -82.7383
0.6586 | -82.7496
0.65891 | -82.755
0.65922 | -82.7548
0.65953 | -82.7611
0.65984 | -82.7701
0.66014 | -82.7876
0.66045 | -82.7988
0.66076 | -82.8064
0.66107 | -82.8093
0.66138 | -82.8112
0.66169 | -82.8153
0.662 | -82.8202
0.66231 | -82.8251
0.66262 | -82.8296
0.66292 | -82.8359
0.66323 | -82.8465
0.66354 | -82.8567
0.66385 | -82.8632
0.66416 | -82.861
0.66447 | -82.8551
0.66478 | -82.8572
0.66509 | -82.8588
0.66539 | -82.8623
0.6657 | -82.8643
0.66601 | -82.8613
0.66632 | -82.8528
0.66663 | -82.8538
0.66694 | -82.8595
0.66725 | -82.8598
0.66756 | -82.8601
0.66787 | -82.8632
0.66817 | -82.8676
0.66848 | -82.8659
0.66879 | -82.8732
0.6691 | -82.8732
0.66941 | -82.8722
0.66972 | -82.873
0.67003 | -82.8753
0.67034 | -82.8735
0.67064 | -82.8709
0.67095 | -82.8736
0.67126 | -82.8739
0.67157 | -82.8775
0.67188 | -82.8789
0.67219 | -82.8788
0.6725 | -82.8824
0.6728 | -82.8872
0.67311 | -82.8878
0.67342 | -82.8884
0.67373 | -82.8906
0.67404 | -82.887
0.67435 | -82.8826
0.67466 | -82.8851
0.67497 | -82.8897
0.67527 | -82.8886
0.67558 | -82.8871
0.67589 | -82.8821
0.6762 | -82.8889
0.67651 | -82.8894
0.67682 | -82.8828
0.67713 | -82.8799
0.67744 | -82.8777
0.67775 | -82.8766
0.67805 | -82.8776
0.67836 | -82.8796
0.67867 | -82.8794
0.67898 | -82.8782
0.67929 | -82.8798
0.6796 | -82.8836
0.67991 | -82.8829
0.68022 | -82.8845
0.68052 | -82.8834
0.68083 | -82.8831
0.68114 | -82.8838
0.68145 | -82.8856
0.68176 | -82.8924
0.68207 | -82.8962
0.68238 | -82.9058
0.68269 | -82.9125
0.68299 | -82.9147
0.6833 | -82.9156
0.68361 | -82.9169
0.68392 | -82.9218
0.68423 | -82.9291
0.68454 | -82.9406
0.68485 | -82.9561
0.68516 | -82.9641
0.68546 | -82.9698
0.68577 | -82.9883
0.68608 | -82.9915
0.68639 | -83.0002
0.6867 | -83.0103
0.68701 | -83.0238
0.68732 | -83.0338
0.68763 | -83.0445
0.68793 | -83.0586
0.68824 | -83.0699
0.68855 | -83.0817
0.68886 | -83.0956
0.68917 | -83.1073
0.68948 | -83.1202
0.68979 | -83.1358
0.6901 | -83.1535
0.69041 | -83.1705
0.69071 | -83.1891
0.69102 | -83.2032
0.69133 | -83.2176
0.69164 | -83.2379
0.69195 | -83.2631
0.69226 | -83.2776
0.69257 | -83.2944
0.69288 | -83.3095
0.69318 | -83.3363
0.69349 | -83.3512
0.6938 | -83.3707
0.69411 | -83.397
0.69442 | -83.4198
0.69473 | -83.437
0.69504 | -83.4562
0.69534 | -83.4749
0.69565 | -83.492
0.69596 | -83.5154
0.69627 | -83.5354
0.69658 | -83.5507
0.69689 | -83.5665
0.6972 | -83.582
0.69751 | -83.6037
0.69781 | -83.62
0.69812 | -83.6286
0.69843 | -83.6459
0.69874 | -83.6694
0.69905 | -83.683
0.69936 | -83.6965
0.69967 | -83.7115
0.69998 | -83.7254
0.70029 | -83.7431
0.70059 | -83.756
0.7009 | -83.7696
0.70121 | -83.7837
0.70152 | -83.797
0.70183 | -83.8119
0.70214 | -83.8329
0.70245 | -83.8401
0.70276 | -83.8404
0.70306 | -83.8502
0.70337 | -83.8593
0.70368 | -83.8626
0.70399 | -83.8711
0.7043 | -83.8818
0.70461 | -83.8928
0.70492 | -83.9037
0.70523 | -83.9133
0.70553 | -83.9148
0.70584 | -83.9131
0.70615 | -83.9191
0.70646 | -83.9251
0.70677 | -83.936
0.70708 | -83.9454
0.70739 | -83.9565
0.7077 | -83.9645
0.708 | -83.9665
0.70831 | -83.9677
0.70862 | -83.9737
0.70893 | -83.9818
0.70924 | -83.9884
0.70955 | -83.996
0.70986 | -84.0037
0.71017 | -84.0066
0.71047 | -84.009
0.71078 | -84.0142
0.71109 | -84.0128
0.7114 | -84.0037
0.71171 | -84.004
0.71202 | -84.0043
0.71233 | -84.0092
0.71264 | -83.9967
0.71295 | -84.0072
0.71325 | -84.002
0.71356 | -83.997
0.71387 | -83.9878
0.71418 | -83.9804
0.71449 | -83.9809
0.7148 | -83.9883
0.71511 | -83.9829
0.71542 | -83.9763
0.71572 | -83.9685
0.71603 | -83.9637
0.71634 | -83.9579
0.71665 | -83.9527
0.71696 | -83.95
0.71727 | -83.9439
0.71758 | -83.9374
0.71788 | -83.9359
0.71819 | -83.9329
0.7185 | -83.9303
0.71881 | -83.9251
0.71912 | -83.918
0.71943 | -83.9079
0.71974 | -83.904
0.72005 | -83.8978
0.72035 | -83.8945
0.72066 | -83.8899
0.72097 | -83.8871
0.72128 | -83.8804
0.72159 | -83.8777
0.7219 | -83.8728
0.72221 | -83.8754
0.72252 | -83.8835
0.72283 | -83.8831
0.72313 | -83.881
0.72344 | -83.8778
0.72375 | -83.872
0.72406 | -83.8582
0.72437 | -83.8505
0.72468 | -83.8457
0.72499 | -83.8432
0.7253 | -83.8369
0.7256 | -83.8301
0.72591 | -83.8324
0.72622 | -83.8364
0.72653 | -83.8331
0.72684 | -83.8312
0.72715 | -83.8287
0.72746 | -83.8267
0.72777 | -83.8248
0.72807 | -83.8221
0.72838 | -83.8151
0.72869 | -83.8175
0.729 | -83.8186
0.72931 | -83.8176
0.72962 | -83.8162
0.72993 | -83.8193
0.73024 | -83.8273
0.73054 | -83.8301
0.73085 | -83.833
0.73116 | -83.8338
0.73147 | -83.8335
0.73178 | -83.8391
0.73209 | -83.8421
0.7324 | -83.8425
0.73271 | -83.8407
0.73301 | -83.8426
0.73332 | -83.8496
0.73363 | -83.8503
0.73394 | -83.8483
0.73425 | -83.8419
0.73456 | -83.8429
0.73487 | -83.8541
0.73518 | -83.8573
0.73549 | -83.8598
0.73579 | -83.8629
0.7361 | -83.8742
0.73641 | -83.881
0.73672 | -83.8778
0.73703 | -83.8773
0.73734 | -83.885
0.73765 | -83.8921
0.73796 | -83.8962
0.73826 | -83.8978
0.73857 | -83.8981
0.73888 | -83.9
0.73919 | -83.9052
0.7395 | -83.9084
0.73981 | -83.9121
0.74012 | -83.915
0.74042 | -83.9173
0.74073 | -83.9197
0.74104 | -83.9211
0.74135 | -83.9233
0.74166 | -83.9271
0.74197 | -83.9303
0.74228 | -83.9339
0.74259 | -83.9418
0.74289 | -83.9475
0.7432 | -83.9539
0.74351 | -83.9575
0.74382 | -83.9636
0.74413 | -83.9674
0.74444 | -83.9691
0.74475 | -83.9713
0.74506 | -83.9692
0.74537 | -83.9693
0.74567 | -83.9679
0.74598 | -83.9724
0.74629 | -83.9764
0.7466 | -83.9763
0.74691 | -83.9747
0.74722 | -83.9732
0.74753 | -83.9758
0.74784 | -83.9762
0.74814 | -83.9782
0.74845 | -83.9704
0.74876 | -83.9563
0.74907 | -83.9606
0.74938 | -83.9615
0.74969 | -83.951
0.75 | -83.9474
0.75031 | -83.9489
0.75061 | -83.9463
0.75092 | -83.9418
0.75123 | -83.9392
0.75154 | -83.9372
0.75185 | -83.9357
0.75216 | -83.9361
0.75247 | -83.9347
0.75278 | -83.9342
0.75309 | -83.9306
0.75339 | -83.9326
0.7537 | -83.9365
0.75401 | -83.9337
0.75432 | -83.9303
0.75463 | -83.9205
0.75494 | -83.9154
0.75525 | -83.9155
0.75555 | -83.9115
0.75586 | -83.9053
0.75617 | -83.9011
0.75648 | -83.8972
0.75679 | -83.8993
0.7571 | -83.8959
0.75741 | -83.8969
0.75772 | -83.8909
0.75802 | -83.8812
0.75833 | -83.8785
0.75864 | -83.8763
0.75895 | -83.8672
0.75926 | -83.8594
0.75957 | -83.8448
0.75988 | -83.8305
0.76019 | -83.8243
0.7605 | -83.8196
0.7608 | -83.8179
0.76111 | -83.8149
0.76142 | -83.8106
0.76173 | -83.8087
0.76204 | -83.8031
0.76235 | -83.7937
0.76266 | -83.7931
0.76297 | -83.8014
0.76327 | -83.8026
0.76358 | -83.7976
0.76389 | -83.7978
0.7642 | -83.8008
0.76451 | -83.7978
0.76482 | -83.7883
0.76513 | -83.7824
0.76543 | -83.7766
0.76574 | -83.7747
0.76605 | -83.77
0.76636 | -83.7664
0.76667 | -83.7598
0.76698 | -83.752
0.76729 | -83.7505
0.7676 | -83.7478
0.7679 | -83.7481
0.76821 | -83.7448
0.76852 | -83.7408
0.76883 | -83.7366
0.76914 | -83.7332
0.76945 | -83.7307
0.76976 | -83.7291
0.77007 | -83.7244
0.77038 | -83.7169
0.77068 | -83.7145
0.77099 | -83.7108
0.7713 | -83.7135
0.77161 | -83.7189
0.77192 | -83.7152
0.77223 | -83.7146
0.77254 | -83.7115
0.77285 | -83.7115
0.77315 | -83.7089
0.77346 | -83.7115
0.77377 | -83.7113
0.77408 | -83.7071
0.77439 | -83.7072
0.7747 | -83.7002
0.77501 | -83.7027
0.77532 | -83.7038
0.77563 | -83.7042
0.77593 | -83.7045
0.77624 | -83.7078
0.77655 | -83.712
0.77686 | -83.7221
0.77717 | -83.727
0.77748 | -83.7309
0.77779 | -83.7331
0.77809 | -83.7347
0.7784 | -83.738
0.77871 | -83.7393
0.77902 | -83.7424
0.77933 | -83.7429
0.77964 | -83.7457
0.77995 | -83.7494
0.78026 | -83.7528
0.78056 | -83.7579
0.78087 | -83.7645
0.78118 | -83.7726
0.78149 | -83.7721
0.7818 | -83.7757
0.78211 | -83.7834
0.78242 | -83.7897
0.78273 | -83.7948
0.78304 | -83.8009
0.78334 | -83.8069
0.78365 | -83.8133
0.78396 | -83.8292
0.78427 | -83.8371
0.78458 | -83.8473
0.78489 | -83.8569
0.7852 | -83.8672
0.78551 | -83.877
0.78581 | -83.889
0.78612 | -83.8996
0.78643 | -83.9135
0.78674 | -83.9258
0.78705 | -83.9393
0.78736 | -83.9544
0.78767 | -83.9736
0.78797 | -83.9879
0.78828 | -83.9998
0.78859 | -84.0104
0.7889 | -84.0178
0.78921 | -84.0292
0.78952 | -84.0434
0.78983 | -84.0564
0.79014 | -84.0762
0.79044 | -84.0945
0.79075 | -84.1056
0.79106 | -84.1193
0.79137 | -84.1343
0.79168 | -84.1495
0.79199 | -84.1689
0.7923 | -84.1898
0.79261 | -84.2154
0.79292 | -84.2312
0.79322 | -84.2507
0.79353 | -84.2716
0.79384 | -84.2883
0.79415 | -84.3069
0.79446 | -84.3219
0.79477 | -84.3425
0.79508 | -84.3653
0.79539 | -84.3893
0.7957 | -84.4124
0.796 | -84.4347
0.79631 | -84.4548
0.79662 | -84.47
0.79693 | -84.4915
0.79724 | -84.5141
0.79755 | -84.5393
0.79786 | -84.5617
0.79817 | -84.5776
0.79847 | -84.6014
0.79878 | -84.6264
0.79909 | -84.6522
0.7994 | -84.6839
0.79971 | -84.7126
0.80002 | -84.734
0.80033 | -84.7526
0.80063 | -84.7696
0.80094 | -84.7806
0.80125 | -84.798
0.80156 | -84.8052
0.80187 | -84.8255
0.80218 | -84.8466
0.80249 | -84.8655
0.8028 | -84.8837
0.8031 | -84.9073
0.80341 | -84.9435
0.80372 | -84.9622
0.80403 | -84.9753
0.80434 | -84.9986
0.80465 | -85.0203
0.80496 | -85.0285
0.80527 | -85.0393
0.80558 | -85.0525
0.80588 | -85.0615
0.80619 | -85.0712
0.8065 | -85.0795
0.80681 | -85.0861
0.80712 | -85.0999
0.80743 | -85.1117
0.80774 | -85.1208
0.80805 | -85.1276
0.80835 | -85.1372
0.80866 | -85.146
0.80897 | -85.1549
0.80928 | -85.1646
0.80959 | -85.1618
0.8099 | -85.1514
0.81021 | -85.1484
0.81051 | -85.1337
0.81082 | -85.1359
0.81113 | -85.1355
0.81144 | -85.1318
0.81175 | -85.1335
0.81206 | -85.1322
0.81237 | -85.1255
0.81268 | -85.1254
0.81299 | -85.1237
0.81329 | -85.1215
0.8136 | -85.1158
0.81391 | -85.1103
0.81422 | -85.1082
0.81453 | -85.1039
0.81484 | -85.11
0.81515 | -85.1136
0.81546 | -85.1023
0.81576 | -85.091
0.81607 | -85.0785
0.81638 | -85.0685
0.81669 | -85.0597
0.817 | -85.0481
0.81731 | -85.039
0.81762 | -85.0286
0.81793 | -85.0078
0.81824 | -84.9918
0.81854 | -84.9795
0.81885 | -84.9672
0.81916 | -84.9521
0.81947 | -84.9373
0.81978 | -84.9179
0.82009 | -84.9006
0.8204 | -84.887
0.82071 | -84.8656
0.82101 | -84.8486
0.82132 | -84.835
0.82163 | -84.8246
0.82194 | -84.8088
0.82225 | -84.7944
0.82256 | -84.7812
0.82287 | -84.7675
0.82317 | -84.7398
0.82348 | -84.709
0.82379 | -84.6899
0.8241 | -84.6751
0.82441 | -84.6596
0.82472 | -84.6435
0.82503 | -84.6276
0.82534 | -84.6127
0.82564 | -84.5956
0.82595 | -84.5828
0.82626 | -84.5711
0.82657 | -84.5522
0.82688 | -84.5398
0.82719 | -84.5237
0.8275 | -84.5047
0.82781 | -84.4888
0.82812 | -84.4762
0.82842 | -84.4704
0.82873 | -84.4584
0.82904 | -84.4445
0.82935 | -84.4357
0.82966 | -84.4245
0.82997 | -84.4168
0.83028 | -84.4092
0.83059 | -84.3985
0.83089 | -84.3923
0.8312 | -84.3835
0.83151 | -84.3739
0.83182 | -84.3638
0.83213 | -84.3519
0.83244 | -84.3429
0.83275 | -84.3358
0.83305 | -84.3286
0.83336 | -84.3166
0.83367 | -84.308
0.83398 | -84.303
0.83429 | -84.2993
0.8346 | -84.2986
0.83491 | -84.2967
0.83522 | -84.2945
0.83553 | -84.2925
0.83583 | -84.2942
0.83614 | -84.2938
0.83645 | -84.2937
0.83676 | -84.2921
0.83707 | -84.2909
0.83738 | -84.3018
0.83769 | -84.302
0.838 | -84.3084
0.8383 | -84.3166
0.83861 | -84.3224
0.83892 | -84.3268
0.83923 | -84.3338
0.83954 | -84.3398
0.83985 | -84.3463
0.84016 | -84.352
0.84047 | -84.3599
0.84078 | -84.3682
0.84108 | -84.3748
0.84139 | -84.3831
0.8417 | -84.3939
0.84201 | -84.407
0.84232 | -84.4169
0.84263 | -84.4205
0.84294 | -84.4226
0.84325 | -84.4418
0.84355 | -84.4536
0.84386 | -84.4639
0.84417 | -84.4757
0.84448 | -84.4902
0.84479 | -84.5052
0.8451 | -84.5199
0.84541 | -84.5349
0.84571 | -84.5554
0.84602 | -84.5787
0.84633 | -84.5995
0.84664 | -84.619
0.84695 | -84.6337
0.84726 | -84.6524
0.84757 | -84.6736
0.84788 | -84.6942
0.84818 | -84.7163
0.84849 | -84.7346
0.8488 | -84.7541
0.84911 | -84.7744
0.84942 | -84.7938
0.84973 | -84.8121
0.85004 | -84.8297
0.85035 | -84.8483
0.85066 | -84.8697
0.85096 | -84.8915
0.85127 | -84.9111
0.85158 | -84.9338
0.85189 | -84.9545
0.8522 | -84.9809
0.85251 | -85.0028
0.85282 | -85.0233
0.85313 | -85.0446
0.85343 | -85.0644
0.85374 | -85.0866
0.85405 | -85.1098
0.85436 | -85.1348
0.85467 | -85.1555
0.85498 | -85.1789
0.85529 | -85.1977
0.85559 | -85.2148
0.8559 | -85.2371
0.85621 | -85.2622
0.85652 | -85.281
0.85683 | -85.3016
0.85714 | -85.3211
0.85745 | -85.3399
0.85776 | -85.3599
0.85807 | -85.3809
0.85837 | -85.4032
0.85868 | -85.428
0.85899 | -85.4584
0.8593 | -85.4826
0.85961 | -85.5103
0.85992 | -85.5313
0.86023 | -85.5455
0.86054 | -85.5603
0.86084 | -85.5738
0.86115 | -85.5866
0.86146 | -85.6073
0.86177 | -85.6244
0.86208 | -85.6421
0.86239 | -85.6584
0.8627 | -85.6743
0.86301 | -85.6912
0.86332 | -85.7143
0.86362 | -85.7325
0.86393 | -85.7462
0.86424 | -85.7608
0.86455 | -85.7762
0.86486 | -85.7904
0.86517 | -85.8057
0.86548 | -85.8239
0.86579 | -85.838
0.86609 | -85.8511
0.8664 | -85.8601
0.86671 | -85.8715
0.86702 | -85.8845
0.86733 | -85.8939
0.86764 | -85.9043
0.86795 | -85.9179
0.86825 | -85.9418
0.86856 | -85.9657
0.86887 | -85.9805
0.86918 | -85.9949
0.86949 | -86.009
0.8698 | -86.0166
0.87011 | -86.0253
0.87042 | -86.0368
0.87072 | -86.0451
0.87103 | -86.0567
0.87134 | -86.0669
0.87165 | -86.0632
0.87196 | -86.0688
0.87227 | -86.0787
0.87258 | -86.0859
0.87289 | -86.0868
0.8732 | -86.0939
0.8735 | -86.0991
0.87381 | -86.1051
0.87412 | -86.1113
0.87443 | -86.1153
0.87474 | -86.1193
0.87505 | -86.1202
0.87536 | -86.1212
0.87567 | -86.1178
0.87597 | -86.1138
0.87628 | -86.1125
0.87659 | -86.1102
0.8769 | -86.1119
0.87721 | -86.1194
0.87752 | -86.1199
0.87783 | -86.1197
0.87814 | -86.1227
0.87844 | -86.1237
0.87875 | -86.1238
0.87906 | -86.1237
0.87937 | -86.1276
0.87968 | -86.1309
0.87999 | -86.1367
0.8803 | -86.1477
0.88061 | -86.1485
0.88091 | -86.1463
0.88122 | -86.1457
0.88153 | -86.1436
0.88184 | -86.1403
0.88215 | -86.1319
0.88246 | -86.1287
0.88277 | -86.1283
0.88308 | -86.1375
0.88338 | -86.1517
0.88369 | -86.15
0.884 | -86.1367
0.88431 | -86.1192
0.88462 | -86.1147
0.88493 | -86.1105
0.88524 | -86.1087
0.88555 | -86.1068
0.88585 | -86.1001
0.88616 | -86.0915
0.88647 | -86.0759
0.88678 | -86.0658
0.88709 | -86.061
0.8874 | -86.0592
0.88771 | -86.0566
0.88802 | -86.0447
0.88833 | -86.0342
0.88863 | -86.0249
0.88894 | -86.0149
0.88925 | -86.0035
0.88956 | -85.993
0.88987 | -85.9832
0.89018 | -85.9698
0.89049 | -85.9517
0.8908 | -85.94
0.8911 | -85.9311
0.89141 | -85.9183
0.89172 | -85.9003
0.89203 | -85.8887
0.89234 | -85.8803
0.89265 | -85.8743
0.89296 | -85.8651
0.89326 | -85.8544
0.89357 | -85.845
0.89388 | -85.8428
0.89419 | -85.8359
0.8945 | -85.8216
0.89481 | -85.8102
0.89512 | -85.7994
0.89543 | -85.7887
0.89573 | -85.7769
0.89604 | -85.764
0.89635 | -85.7537
0.89666 | -85.743
0.89697 | -85.7314
0.89728 | -85.7215
0.89759 | -85.7096
0.8979 | -85.7025
0.89821 | -85.6822
0.89851 | -85.6605
0.89882 | -85.6427
0.89913 | -85.6301
0.89944 | -85.6194
0.89975 | -85.6088
0.90006 | -85.5931
0.90037 | -85.5801
0.90068 | -85.5681
0.90098 | -85.5565
0.90129 | -85.5468
0.9016 | -85.5324
0.90191 | -85.5222
0.90222 | -85.5138
0.90253 | -85.5047
0.90284 | -85.4922
0.90315 | -85.4803
0.90345 | -85.4704
0.90376 | -85.4658
0.90407 | -85.4639
0.90438 | -85.4515
0.90469 | -85.4431
0.905 | -85.4343
0.90531 | -85.4261
0.90562 | -85.4162
0.90592 | -85.4061
0.90623 | -85.3955
0.90654 | -85.3847
0.90685 | -85.3726
0.90716 | -85.3639
0.90747 | -85.3548
0.90778 | -85.3468
0.90809 | -85.3376
0.90839 | -85.324
0.9087 | -85.3138
0.90901 | -85.3038
0.90932 | -85.2948
0.90963 | -85.2776
0.90994 | -85.2597
0.91025 | -85.2475
0.91056 | -85.238
0.91087 | -85.2303
0.91117 | -85.2201
0.91148 | -85.2062
0.91179 | -85.1986
0.9121 | -85.1925
0.91241 | -85.1859
0.91272 | -85.1782
0.91303 | -85.174
0.91334 | -85.1685
0.91364 | -85.1638
0.91395 | -85.1577
0.91426 | -85.1515
0.91457 | -85.1478
0.91488 | -85.1436
0.91519 | -85.1386
0.9155 | -85.1314
0.9158 | -85.1252
0.91611 | -85.1238
0.91642 | -85.1229
0.91673 | -85.1201
0.91704 | -85.1167
0.91735 | -85.1194
0.91766 | -85.12
0.91797 | -85.1221
0.91827 | -85.1175
0.91858 | -85.1161
0.91889 | -85.1168
0.9192 | -85.1113
0.91951 | -85.111
0.91982 | -85.1037
0.92013 | -85.0869
0.92044 | -85.0777
0.92075 | -85.0786
0.92105 | -85.0785
0.92136 | -85.078
0.92167 | -85.0794
0.92198 | -85.0777
0.92229 | -85.0764
0.9226 | -85.077
0.92291 | -85.078
0.92322 | -85.0815
0.92352 | -85.0849
0.92383 | -85.0848
0.92414 | -85.083
0.92445 | -85.0823
0.92476 | -85.0842
0.92507 | -85.0832
0.92538 | -85.0842
0.92569 | -85.0844
0.926 | -85.0845
0.9263 | -85.0904
0.92661 | -85.092
0.92692 | -85.0946
0.92723 | -85.0978
0.92754 | -85.1021
0.92785 | -85.1058
0.92816 | -85.1068
0.92846 | -85.1087
0.92877 | -85.109
0.92908 | -85.1093
0.92939 | -85.1138
0.9297 | -85.1163
0.93001 | -85.1147
0.93032 | -85.1151
0.93063 | -85.1183
0.93093 | -85.1252
0.93124 | -85.1301
0.93155 | -85.1336
0.93186 | -85.1375
0.93217 | -85.1405
0.93248 | -85.1433
0.93279 | -85.1489
0.9331 | -85.1517
0.93341 | -85.1533
0.93371 | -85.152
0.93402 | -85.1546
0.93433 | -85.1625
0.93464 | -85.1727
0.93495 | -85.1797
0.93526 | -85.1828
0.93557 | -85.187
0.93588 | -85.1911
0.93618 | -85.1942
0.93649 | -85.1961
0.9368 | -85.199
0.93711 | -85.2047
0.93742 | -85.2084
0.93773 | -85.2084
0.93804 | -85.2123
0.93834 | -85.2157
0.93865 | -85.2204
0.93896 | -85.2277
0.93927 | -85.2373
0.93958 | -85.2417
0.93989 | -85.24
0.9402 | -85.2419
0.94051 | -85.2493
0.94081 | -85.254
0.94112 | -85.2584
0.94143 | -85.2643
0.94174 | -85.2704
0.94205 | -85.2756
0.94236 | -85.2796
0.94267 | -85.2836
0.94298 | -85.287
0.94329 | -85.2919
0.94359 | -85.2943
0.9439 | -85.294
0.94421 | -85.2945
0.94452 | -85.2981
0.94483 | -85.2994
0.94514 | -85.301
0.94545 | -85.3053
0.94576 | -85.3111
0.94606 | -85.3164
0.94637 | -85.3255
0.94668 | -85.33
0.94699 | -85.3335
0.9473 | -85.3397
0.94761 | -85.3452
0.94792 | -85.3507
0.94823 | -85.3627
0.94854 | -85.3766
0.94884 | -85.3829
0.94915 | -85.3881
0.94946 | -85.3913
0.94977 | -85.3953
0.95008 | -85.4004
0.95039 | -85.4104
0.9507 | -85.4276
0.951 | -85.432
0.95131 | -85.4344
0.95162 | -85.4386
0.95193 | -85.4423
0.95224 | -85.4503
0.95255 | -85.456
0.95286 | -85.4604
0.95317 | -85.4705
0.95347 | -85.4775
0.95378 | -85.482
0.95409 | -85.4879
0.9544 | -85.4942
0.95471 | -85.5035
0.95502 | -85.5156
0.95533 | -85.5228
0.95564 | -85.5284
0.95595 | -85.5346
0.95625 | -85.5467
0.95656 | -85.5611
0.95687 | -85.5685
0.95718 | -85.5749
0.95749 | -85.5801
0.9578 | -85.5827
0.95811 | -85.5854
0.95842 | -85.5808
0.95872 | -85.5776
0.95903 | -85.581
0.95934 | -85.5876
0.95965 | -85.5918
0.95996 | -85.5973
0.96027 | -85.6008
0.96058 | -85.605
0.96088 | -85.6117
0.96119 | -85.6179
0.9615 | -85.6255
0.96181 | -85.6326
0.96212 | -85.6367
0.96243 | -85.6442
0.96274 | -85.6549
0.96305 | -85.6614
0.96335 | -85.6773
0.96366 | -85.6974
0.96397 | -85.7085
0.96428 | -85.7138
0.96459 | -85.7132
0.9649 | -85.7112
0.96521 | -85.7142
0.96552 | -85.7254
0.96583 | -85.7252
0.96613 | -85.6988
0.96644 | -85.6653
0.96675 | -85.6643
0.96706 | -85.6676
0.96737 | -85.6639
0.96768 | -85.6656
0.96799 | -85.6644
0.9683 | -85.6658
0.9686 | -85.66
0.96891 | -85.6511
0.96922 | -85.6514
0.96953 | -85.6539
0.96984 | -85.6549
0.97015 | -85.6544
0.97046 | -85.6525
0.97077 | -85.6484
0.97108 | -85.6443
0.97138 | -85.6378
0.97169 | -85.6324
0.972 | -85.628
0.97231 | -85.6218
0.97262 | -85.6175
0.97293 | -85.6117
0.97324 | -85.6055
0.97354 | -85.6024
0.97385 | -85.5995
0.97416 | -85.6013
0.97447 | -85.5955
0.97478 | -85.5917
0.97509 | -85.5883
0.9754 | -85.5819
0.97571 | -85.583
0.97601 | -85.5851
0.97632 | -85.5809
0.97663 | -85.574
0.97694 | -85.5654
0.97725 | -85.5584
0.97756 | -85.5559
0.97787 | -85.5474
0.97818 | -85.5365
0.97849 | -85.5259
0.97879 | -85.5153
0.9791 | -85.5042
0.97941 | -85.4925
0.97972 | -85.4798
0.98003 | -85.469
0.98034 | -85.4539
0.98065 | -85.4294
0.98096 | -85.4136
0.98126 | -85.4025
0.98157 | -85.3927
0.98188 | -85.3832
0.98219 | -85.3732
0.9825 | -85.3616
0.98281 | -85.3484
0.98312 | -85.3331
0.98342 | -85.318
0.98373 | -85.3036
0.98404 | -85.2912
0.98435 | -85.2783
0.98466 | -85.266
0.98497 | -85.2566
0.98528 | -85.2449
0.98559 | -85.232
0.98589 | -85.2232
0.9862 | -85.2102
0.98651 | -85.2021
0.98682 | -85.1998
0.98713 | -85.1874
0.98744 | -85.1708
0.98775 | -85.1562
0.98806 | -85.1435
0.98837 | -85.1302
0.98867 | -85.1197
0.98898 | -85.1051
0.98929 | -85.0923
0.9896 | -85.0813
0.98991 | -85.0688
0.99022 | -85.0564
0.99053 | -85.0443
0.99084 | -85.034
0.99115 | -85.0245
0.99145 | -85.0128
0.99176 | -85.0019
0.99207 | -84.9924
0.99238 | -84.9829
0.99269 | -84.9743
0.993 | -84.9667
0.99331 | -84.9571
0.99362 | -84.9484
0.99392 | -84.9369
0.99423 | -84.9244
0.99454 | -84.9169
0.99485 | -84.9099
0.99516 | -84.9037
0.99547 | -84.8984
0.99578 | -84.892
0.99608 | -84.8855
0.99639 | -84.8831
0.9967 | -84.8816
0.99701 | -84.8826
0.99732 | -84.8836
0.99763 | -84.8928
0.99794 | -84.8943
0.99825 | -84.8931
0.99855 | -84.891
0.99886 | -84.8909
0.99917 | -84.8919
0.99948 | -84.8934
0.99979 | -84.8942
1.0001 | -84.8946
1.0004 | -84.8955
1.0007 | -84.8976
1.001 | -84.9023
1.0013 | -84.9091
1.0016 | -84.9111
1.002 | -84.9092
1.0023 | -84.9077
1.0026 | -84.9071
1.0029 | -84.9118
1.0032 | -84.9185
1.0035 | -84.9231
1.0038 | -84.9279
1.0041 | -84.9344
1.0044 | -84.9464
1.0047 | -84.9549
1.005 | -84.9629
1.0053 | -84.9728
1.0057 | -84.9851
1.006 | -84.9963
1.0063 | -85.0062
1.0066 | -85.0191
1.0069 | -85.0291
1.0072 | -85.0329
1.0075 | -85.0391
1.0078 | -85.0503
1.0081 | -85.0665
1.0084 | -85.0793
1.0087 | -85.0874
1.0091 | -85.0987
1.0094 | -85.11
1.0097 | -85.1195
1.01 | -85.1309
1.0103 | -85.1413
1.0106 | -85.1529
1.0109 | -85.1639
1.0112 | -85.1768
1.0115 | -85.1892
1.0118 | -85.2018
1.0121 | -85.2138
1.0125 | -85.2282
1.0128 | -85.2423
1.0131 | -85.256
1.0134 | -85.2701
1.0137 | -85.2847
1.014 | -85.2967
1.0143 | -85.3091
1.0146 | -85.3217
1.0149 | -85.3341
1.0152 | -85.3457
1.0155 | -85.3534
1.0158 | -85.3626
1.0162 | -85.3711
1.0165 | -85.3808
1.0168 | -85.3892
1.0171 | -85.3977
1.0174 | -85.4089
1.0177 | -85.4194
1.018 | -85.4272
1.0183 | -85.4342
1.0186 | -85.4412
1.0189 | -85.4502
1.0192 | -85.456
1.0196 | -85.4565
1.0199 | -85.4628
1.0202 | -85.4741
1.0205 | -85.4835
1.0208 | -85.4924
1.0211 | -85.5007
1.0214 | -85.508
1.0217 | -85.5156
1.022 | -85.5222
1.0223 | -85.5279
1.0226 | -85.5333
1.0229 | -85.5416
1.0233 | -85.5487
1.0236 | -85.5541
1.0239 | -85.5589
1.0242 | -85.5643
1.0245 | -85.5651
1.0248 | -85.5708
1.0251 | -85.5749
1.0254 | -85.5793
1.0257 | -85.5812
1.026 | -85.5818
1.0263 | -85.5851
1.0267 | -85.5884
1.027 | -85.5932
1.0273 | -85.5988
1.0276 | -85.6039
1.0279 | -85.6149
1.0282 | -85.6235
1.0285 | -85.627
1.0288 | -85.6291
1.0291 | -85.631
1.0294 | -85.6305
1.0297 | -85.6303
1.03 | -85.625
1.0304 | -85.6237
1.0307 | -85.6241
1.031 | -85.624
1.0313 | -85.6208
1.0316 | -85.6165
1.0319 | -85.6145
1.0322 | -85.6113
1.0325 | -85.6085
1.0328 | -85.6065
1.0331 | -85.6026
1.0334 | -85.5973
1.0338 | -85.5959
1.0341 | -85.5951
1.0344 | -85.5954
1.0347 | -85.5986
1.035 | -85.5973
1.0353 | -85.5932
1.0356 | -85.587
1.0359 | -85.5858
1.0362 | -85.6077
1.0365 | -85.637
1.0368 | -85.6353
1.0372 | -85.6312
1.0375 | -85.6266
1.0378 | -85.6177
1.0381 | -85.6123
1.0384 | -85.6052
1.0387 | -85.6031
1.039 | -85.6051
1.0393 | -85.6022
1.0396 | -85.5966
1.0399 | -85.5922
1.0402 | -85.5852
1.0405 | -85.5779
1.0409 | -85.5726
1.0412 | -85.5687
1.0415 | -85.566
1.0418 | -85.5652
1.0421 | -85.5521
1.0424 | -85.5467
1.0427 | -85.5432
1.043 | -85.5424
1.0433 | -85.5362
1.0436 | -85.5297
1.0439 | -85.527
1.0443 | -85.5227
1.0446 | -85.5166
1.0449 | -85.5104
1.0452 | -85.5027
1.0455 | -85.4952
1.0458 | -85.4897
1.0461 | -85.4824
1.0464 | -85.472
1.0467 | -85.4664
1.047 | -85.462
1.0473 | -85.4575
1.0476 | -85.4532
1.048 | -85.4492
1.0483 | -85.4462
1.0486 | -85.4422
1.0489 | -85.4371
1.0492 | -85.4344
1.0495 | -85.4303
1.0498 | -85.4268
1.0501 | -85.4216
1.0504 | -85.4127
1.0507 | -85.4087
1.051 | -85.4059
1.0514 | -85.3994
1.0517 | -85.3999
1.052 | -85.4091
1.0523 | -85.4079
1.0526 | -85.4014
1.0529 | -85.396
1.0532 | -85.391
1.0535 | -85.3865
1.0538 | -85.3808
1.0541 | -85.3732
1.0544 | -85.3699
1.0548 | -85.368
1.0551 | -85.3625
1.0554 | -85.3563
1.0557 | -85.3522
1.056 | -85.3488
1.0563 | -85.3435
1.0566 | -85.3393
1.0569 | -85.3325
1.0572 | -85.3278
1.0575 | -85.3258
1.0578 | -85.3249
1.0581 | -85.3206
1.0585 | -85.3213
1.0588 | -85.3184
1.0591 | -85.3142
1.0594 | -85.3084
1.0597 | -85.3026
1.06 | -85.3021
1.0603 | -85.2954
1.0606 | -85.2858
1.0609 | -85.2821
1.0612 | -85.279
1.0615 | -85.275
1.0619 | -85.2712
1.0622 | -85.263
1.0625 | -85.2534
1.0628 | -85.2457
1.0631 | -85.2404
1.0634 | -85.2355
1.0637 | -85.2332
1.064 | -85.2308
1.0643 | -85.2261
1.0646 | -85.2227
1.0649 | -85.2201
1.0652 | -85.2162
1.0656 | -85.2115
1.0659 | -85.2103
1.0662 | -85.2121
1.0665 | -85.2107
1.0668 | -85.208
1.0671 | -85.2048
1.0674 | -85.204
1.0677 | -85.2025
1.068 | -85.2033
1.0683 | -85.1999
1.0686 | -85.1962
1.069 | -85.198
1.0693 | -85.1961
1.0696 | -85.1927
1.0699 | -85.1861
1.0702 | -85.181
1.0705 | -85.1784
1.0708 | -85.1776
1.0711 | -85.1771
1.0714 | -85.1763
1.0717 | -85.1758
1.072 | -85.1751
1.0723 | -85.1763
1.0727 | -85.1784
1.073 | -85.1781
1.0733 | -85.1791
1.0736 | -85.1807
1.0739 | -85.1827
1.0742 | -85.1865
1.0745 | -85.1929
1.0748 | -85.2004
1.0751 | -85.2065
1.0754 | -85.2045
1.0757 | -85.1964
1.0761 | -85.1938
1.0764 | -85.1954
1.0767 | -85.1968
1.077 | -85.1947
1.0773 | -85.1952
1.0776 | -85.1964
1.0779 | -85.2044
1.0782 | -85.2036
1.0785 | -85.2046
1.0788 | -85.2057
1.0791 | -85.209
1.0795 | -85.2146
1.0798 | -85.2188
1.0801 | -85.2234
1.0804 | -85.2301
1.0807 | -85.2345
1.081 | -85.2346
1.0813 | -85.2235
1.0816 | -85.2176
1.0819 | -85.2183
1.0822 | -85.2194
1.0825 | -85.222
1.0828 | -85.2225
1.0832 | -85.2185
1.0835 | -85.2177
1.0838 | -85.2199
1.0841 | -85.2232
1.0844 | -85.2263
1.0847 | -85.2276
1.085 | -85.2286
1.0853 | -85.2312
1.0856 | -85.2341
1.0859 | -85.2358
1.0862 | -85.2354
1.0866 | -85.2374
1.0869 | -85.2406
1.0872 | -85.2433
1.0875 | -85.2445
1.0878 | -85.2456
1.0881 | -85.2475
1.0884 | -85.2498
1.0887 | -85.2545
1.089 | -85.2599
1.0893 | -85.264
1.0896 | -85.2698
1.0899 | -85.2748
1.0903 | -85.2783
1.0906 | -85.2831
1.0909 | -85.2862
1.0912 | -85.2873
1.0915 | -85.286
1.0918 | -85.289
1.0921 | -85.2937
1.0924 | -85.2975
1.0927 | -85.3012
1.093 | -85.3059
1.0933 | -85.3144
1.0937 | -85.3159
1.094 | -85.3138
1.0943 | -85.3119
1.0946 | -85.3122
1.0949 | -85.3129
1.0952 | -85.3139
1.0955 | -85.3137
1.0958 | -85.3143
1.0961 | -85.3146
1.0964 | -85.3149
1.0967 | -85.3168
1.0971 | -85.3181
1.0974 | -85.3164
1.0977 | -85.3162
1.098 | -85.3173
1.0983 | -85.3171
1.0986 | -85.318
1.0989 | -85.318
1.0992 | -85.3177
1.0995 | -85.3176
1.0998 | -85.3177
1.1001 | -85.3191
1.1004 | -85.3191
1.1008 | -85.3175
1.1011 | -85.318
1.1014 | -85.3192
1.1017 | -85.3221
1.102 | -85.3204
1.1023 | -85.3181
1.1026 | -85.3187
1.1029 | -85.3182
1.1032 | -85.3179
1.1035 | -85.3192
1.1038 | -85.3218
1.1042 | -85.3249
1.1045 | -85.3277
1.1048 | -85.3358
1.1051 | -85.3389
1.1054 | -85.3396
1.1057 | -85.3404
1.106 | -85.345
1.1063 | -85.351
1.1066 | -85.3558
1.1069 | -85.3606
1.1072 | -85.3633
1.1076 | -85.3616
1.1079 | -85.3615
1.1082 | -85.3643
1.1085 | -85.3693
1.1088 | -85.3705
1.1091 | -85.3714
1.1094 | -85.3726
1.1097 | -85.3752
1.11 | -85.3827
1.1103 | -85.3881
1.1106 | -85.3892
1.1109 | -85.3944
1.1113 | -85.3993
1.1116 | -85.4005
1.1119 | -85.401
1.1122 | -85.4035
1.1125 | -85.4102
1.1128 | -85.417
1.1131 | -85.4226
1.1134 | -85.4327
1.1137 | -85.4398
1.114 | -85.4478
1.1143 | -85.4539
1.1147 | -85.4599
1.115 | -85.4655
1.1153 | -85.4701
1.1156 | -85.4758
1.1159 | -85.4847
1.1162 | -85.4914
1.1165 | -85.4989
1.1168 | -85.5067
1.1171 | -85.5146
1.1174 | -85.5247
1.1177 | -85.5412
1.118 | -85.5517
1.1184 | -85.5604
1.1187 | -85.5696
1.119 | -85.5826
1.1193 | -85.5951
1.1196 | -85.6047
1.1199 | -85.6135
1.1202 | -85.625
1.1205 | -85.6391
1.1208 | -85.6564
1.1211 | -85.669
1.1214 | -85.6792
1.1218 | -85.6931
1.1221 | -85.7121
1.1224 | -85.7319
1.1227 | -85.742
1.123 | -85.749
1.1233 | -85.7579
1.1236 | -85.7704
1.1239 | -85.7839
1.1242 | -85.7937
1.1245 | -85.8062
1.1248 | -85.8205
1.1252 | -85.8319
1.1255 | -85.8452
1.1258 | -85.8592
1.1261 | -85.8763
1.1264 | -85.897
1.1267 | -85.9123
1.127 | -85.9276
1.1273 | -85.9454
1.1276 | -85.9629
1.1279 | -85.9781
1.1282 | -85.9914
1.1285 | -86.004
1.1289 | -86.0225
1.1292 | -86.0392
1.1295 | -86.0571
1.1298 | -86.0757
1.1301 | -86.0921
1.1304 | -86.1135
1.1307 | -86.1321
1.131 | -86.1457
1.1313 | -86.1573
1.1316 | -86.1782
1.1319 | -86.1989
1.1323 | -86.2091
1.1326 | -86.2145
1.1329 | -86.2296
1.1332 | -86.2446
1.1335 | -86.2633
1.1338 | -86.2812
1.1341 | -86.2979
1.1344 | -86.3143
1.1347 | -86.3356
1.135 | -86.3523
1.1353 | -86.3696
1.1356 | -86.3851
1.136 | -86.3989
1.1363 | -86.4153
1.1366 | -86.4317
1.1369 | -86.4482
1.1372 | -86.4679
1.1375 | -86.4888
1.1378 | -86.5016
1.1381 | -86.5109
1.1384 | -86.5225
1.1387 | -86.5381
1.139 | -86.5576
1.1394 | -86.5786
1.1397 | -86.596
1.14 | -86.6108
1.1403 | -86.6091
1.1406 | -86.6186
1.1409 | -86.6246
1.1412 | -86.6339
1.1415 | -86.6448
1.1418 | -86.655
1.1421 | -86.6639
1.1424 | -86.674
1.1427 | -86.6879
1.1431 | -86.698
1.1434 | -86.7
1.1437 | -86.7042
1.144 | -86.7098
1.1443 | -86.7166
1.1446 | -86.7236
1.1449 | -86.7311
1.1452 | -86.7375
1.1455 | -86.7444
1.1458 | -86.7501
1.1461 | -86.7493
1.1465 | -86.7518
1.1468 | -86.7558
1.1471 | -86.7611
1.1474 | -86.7694
1.1477 | -86.777
1.148 | -86.7802
1.1483 | -86.7786
1.1486 | -86.7795
1.1489 | -86.7728
1.1492 | -86.7668
1.1495 | -86.7555
1.1499 | -86.7508
1.1502 | -86.7461
1.1505 | -86.7424
1.1508 | -86.7418
1.1511 | -86.7422
1.1514 | -86.7406
1.1517 | -86.7373
1.152 | -86.734
1.1523 | -86.7319
1.1526 | -86.7287
1.1529 | -86.7242
1.1532 | -86.7257
1.1536 | -86.7327
1.1539 | -86.7329
1.1542 | -86.723
1.1545 | -86.7144
1.1548 | -86.7122
1.1551 | -86.7026
1.1554 | -86.691
1.1557 | -86.6814
1.156 | -86.6706
1.1563 | -86.6651
1.1566 | -86.6572
1.157 | -86.6498
1.1573 | -86.6424
1.1576 | -86.6341
1.1579 | -86.6224
1.1582 | -86.6133
1.1585 | -86.6045
1.1588 | -86.5983
1.1591 | -86.5929
1.1594 | -86.589
1.1597 | -86.5969
1.16 | -86.5948
1.1603 | -86.5819
1.1607 | -86.5705
1.161 | -86.5586
1.1613 | -86.5504
1.1616 | -86.5461
1.1619 | -86.5375
1.1622 | -86.5215
1.1625 | -86.5079
1.1628 | -86.4946
1.1631 | -86.4821
1.1634 | -86.4733
1.1637 | -86.4665
1.1641 | -86.4568
1.1644 | -86.4461
1.1647 | -86.4383
1.165 | -86.4297
1.1653 | -86.4203
1.1656 | -86.4103
1.1659 | -86.3985
1.1662 | -86.3894
1.1665 | -86.381
1.1668 | -86.3789
1.1671 | -86.3707
1.1675 | -86.3618
1.1678 | -86.3504
1.1681 | -86.3387
1.1684 | -86.3281
1.1687 | -86.3202
1.169 | -86.3127
1.1693 | -86.3056
1.1696 | -86.2999
1.1699 | -86.2938
1.1702 | -86.2867
1.1705 | -86.2803
1.1708 | -86.2753
1.1712 | -86.2713
1.1715 | -86.2673
1.1718 | -86.261
1.1721 | -86.255
1.1724 | -86.2494
1.1727 | -86.2469
1.173 | -86.2441
1.1733 | -86.2414
1.1736 | -86.2377
1.1739 | -86.2348
1.1742 | -86.2325
1.1746 | -86.2313
1.1749 | -86.2301
1.1752 | -86.2293
1.1755 | -86.2259
1.1758 | -86.2261
1.1761 | -86.2245
1.1764 | -86.224
1.1767 | -86.2245
1.177 | -86.2274
1.1773 | -86.2289
1.1776 | -86.2291
1.1779 | -86.2314
1.1783 | -86.2343
1.1786 | -86.2366
1.1789 | -86.2401
1.1792 | -86.2436
1.1795 | -86.2471
1.1798 | -86.2521
1.1801 | -86.259
1.1804 | -86.261
1.1807 | -86.2581
1.181 | -86.2603
1.1813 | -86.2643
1.1817 | -86.2742
1.182 | -86.2827
1.1823 | -86.2862
1.1826 | -86.2915
1.1829 | -86.299
1.1832 | -86.3053
1.1835 | -86.3114
1.1838 | -86.3179
1.1841 | -86.3243
1.1844 | -86.3287
1.1847 | -86.3384
1.1851 | -86.3441
1.1854 | -86.3479
1.1857 | -86.3545
1.186 | -86.3604
1.1863 | -86.3631
1.1866 | -86.3713
1.1869 | -86.3798
1.1872 | -86.3881
1.1875 | -86.3973
1.1878 | -86.4106
1.1881 | -86.422
1.1884 | -86.432
1.1888 | -86.4401
1.1891 | -86.4456
1.1894 | -86.4529
1.1897 | -86.4618
1.19 | -86.4683
1.1903 | -86.4774
1.1906 | -86.4879
1.1909 | -86.5012
1.1912 | -86.5114
1.1915 | -86.5182
1.1918 | -86.5265
1.1922 | -86.5373
1.1925 | -86.5475
1.1928 | -86.5575
1.1931 | -86.5679
1.1934 | -86.5782
1.1937 | -86.5849
1.194 | -86.5947
1.1943 | -86.6045
1.1946 | -86.6143
1.1949 | -86.6242
1.1952 | -86.6332
1.1955 | -86.6401
1.1959 | -86.6475
1.1962 | -86.6567
1.1965 | -86.669
1.1968 | -86.6777
1.1971 | -86.6872
1.1974 | -86.6966
1.1977 | -86.7058
1.198 | -86.7159
1.1983 | -86.7247
1.1986 | -86.7307
1.1989 | -86.7383
1.1993 | -86.7457
1.1996 | -86.7522
1.1999 | -86.7607
1.2002 | -86.7693
1.2005 | -86.7777
1.2008 | -86.7868
1.2011 | -86.7949
1.2014 | -86.8029
1.2017 | -86.8111
1.202 | -86.8194
1.2023 | -86.8249
1.2026 | -86.8237
1.203 | -86.8316
1.2033 | -86.837
1.2036 | -86.8438
1.2039 | -86.8513
1.2042 | -86.8618
1.2045 | -86.8752
1.2048 | -86.8847
1.2051 | -86.8893
1.2054 | -86.8905
1.2057 | -86.8952
1.206 | -86.8981
1.2064 | -86.9024
1.2067 | -86.9066
1.207 | -86.9093
1.2073 | -86.913
1.2076 | -86.9161
1.2079 | -86.9183
1.2082 | -86.9158
1.2085 | -86.9096
1.2088 | -86.9057
1.2091 | -86.9088
1.2094 | -86.9127
1.2098 | -86.9179
1.2101 | -86.9229
1.2104 | -86.9238
1.2107 | -86.9259
1.211 | -86.9264
1.2113 | -86.9239
1.2116 | -86.9248
1.2119 | -86.9266
1.2122 | -86.9255
1.2125 | -86.9257
1.2128 | -86.9277
1.2131 | -86.9314
1.2135 | -86.9241
1.2138 | -86.9111
1.2141 | -86.9059
1.2144 | -86.9132
1.2147 | -86.9203
1.215 | -86.9068
1.2153 | -86.8985
1.2156 | -86.8951
1.2159 | -86.8865
1.2162 | -86.8775
1.2165 | -86.8733
1.2169 | -86.867
1.2172 | -86.8625
1.2175 | -86.8592
1.2178 | -86.8563
1.2181 | -86.8541
1.2184 | -86.8506
1.2187 | -86.845
1.219 | -86.8397
1.2193 | -86.8348
1.2196 | -86.8327
1.2199 | -86.8305
1.2203 | -86.8203
1.2206 | -86.8154
1.2209 | -86.8081
1.2212 | -86.7997
1.2215 | -86.7936
1.2218 | -86.788
1.2221 | -86.7807
1.2224 | -86.7756
1.2227 | -86.7704
1.223 | -86.7729
1.2233 | -86.7712
1.2236 | -86.7683
1.224 | -86.7625
1.2243 | -86.7551
1.2246 | -86.7479
1.2249 | -86.74
1.2252 | -86.7339
1.2255 | -86.727
1.2258 | -86.7199
1.2261 | -86.718
1.2264 | -86.7127
1.2267 | -86.7047
1.227 | -86.6955
1.2274 | -86.6857
1.2277 | -86.6781
1.228 | -86.6707
1.2283 | -86.6636
1.2286 | -86.6567
1.2289 | -86.6539
1.2292 | -86.6475
1.2295 | -86.6406
1.2298 | -86.6325
1.2301 | -86.6263
1.2304 | -86.6193
1.2307 | -86.6104
1.2311 | -86.6001
1.2314 | -86.5894
1.2317 | -86.5818
1.232 | -86.5751
1.2323 | -86.567
1.2326 | -86.5653
1.2329 | -86.5678
1.2332 | -86.5641
1.2335 | -86.5597
1.2338 | -86.5537
1.2341 | -86.5481
1.2345 | -86.5411
1.2348 | -86.5322
1.2351 | -86.524
1.2354 | -86.5177
1.2357 | -86.5115
1.236 | -86.5059
1.2363 | -86.4958
1.2366 | -86.4879
1.2369 | -86.4825
1.2372 | -86.4761
1.2375 | -86.47
1.2378 | -86.4652
1.2382 | -86.4631
1.2385 | -86.4574
1.2388 | -86.4534
1.2391 | -86.4512
1.2394 | -86.4441
1.2397 | -86.44
1.24 | -86.4351
1.2403 | -86.4285
1.2406 | -86.4201
1.2409 | -86.4159
1.2412 | -86.4106
1.2416 | -86.4064
1.2419 | -86.403
1.2422 | -86.3992
1.2425 | -86.3954
1.2428 | -86.3901
1.2431 | -86.3843
1.2434 | -86.3818
1.2437 | -86.3713
1.244 | -86.3654
1.2443 | -86.3662
1.2446 | -86.364
1.245 | -86.3568
1.2453 | -86.3471
1.2456 | -86.3405
1.2459 | -86.335
1.2462 | -86.3345
1.2465 | -86.3329
1.2468 | -86.3315
1.2471 | -86.3337
1.2474 | -86.3312
1.2477 | -86.3257
1.248 | -86.3221
1.2483 | -86.3203
1.2487 | -86.3188
1.249 | -86.3165
1.2493 | -86.3146
1.2496 | -86.3129
1.2499 | -86.3122
1.2502 | -86.3127
1.2505 | -86.3128
1.2508 | -86.3089
1.2511 | -86.3061
1.2514 | -86.3061
1.2517 | -86.305
1.2521 | -86.3049
1.2524 | -86.3077
1.2527 | -86.3113
1.253 | -86.3134
1.2533 | -86.3046
1.2536 | -86.296
1.2539 | -86.2953
1.2542 | -86.2954
1.2545 | -86.2973
1.2548 | -86.301
1.2551 | -86.3061
1.2554 | -86.3069
1.2558 | -86.3053
1.2561 | -86.3042
1.2564 | -86.3009
1.2567 | -86.301
1.257 | -86.3034
1.2573 | -86.3067
1.2576 | -86.3112
1.2579 | -86.3129
1.2582 | -86.3112
1.2585 | -86.3089
1.2588 | -86.3075
1.2592 | -86.3048
1.2595 | -86.3063
1.2598 | -86.3082
1.2601 | -86.3087
1.2604 | -86.3116
1.2607 | -86.3176
1.261 | -86.3232
1.2613 | -86.3225
1.2616 | -86.3225
1.2619 | -86.3226
1.2622 | -86.3227
1.2626 | -86.3214
1.2629 | -86.3172
1.2632 | -86.319
1.2635 | -86.3143
1.2638 | -86.3123
1.2641 | -86.3113
1.2644 | -86.313
1.2647 | -86.3142
1.265 | -86.3178
1.2653 | -86.3196
1.2656 | -86.3211
1.2659 | -86.3255
1.2663 | -86.333
1.2666 | -86.3358
1.2669 | -86.3376
1.2672 | -86.3358
1.2675 | -86.3371
1.2678 | -86.3408
1.2681 | -86.3423
1.2684 | -86.3407
1.2687 | -86.3386
1.269 | -86.3382
1.2693 | -86.3387
1.2697 | -86.3391
1.27 | -86.3426
1.2703 | -86.3458
1.2706 | -86.3466
1.2709 | -86.3453
1.2712 | -86.345
1.2715 | -86.3451
1.2718 | -86.3465
1.2721 | -86.3515
1.2724 | -86.3505
1.2727 | -86.3482
1.2731 | -86.3474
1.2734 | -86.3472
1.2737 | -86.3463
1.274 | -86.3464
1.2743 | -86.3432
1.2746 | -86.3319
1.2749 | -86.3259
1.2752 | -86.3276
1.2755 | -86.3266
1.2758 | -86.325
1.2761 | -86.3257
1.2764 | -86.3208
1.2768 | -86.3191
1.2771 | -86.3186
1.2774 | -86.3164
1.2777 | -86.3151
1.278 | -86.314
1.2783 | -86.312
1.2786 | -86.3092
1.2789 | -86.306
1.2792 | -86.3029
1.2795 | -86.2985
1.2798 | -86.2963
1.2802 | -86.2941
1.2805 | -86.2908
1.2808 | -86.2874
1.2811 | -86.2827
1.2814 | -86.2766
1.2817 | -86.273
1.282 | -86.2714
1.2823 | -86.2681
1.2826 | -86.2627
1.2829 | -86.2596
1.2832 | -86.2554
1.2835 | -86.2509
1.2839 | -86.246
1.2842 | -86.2416
1.2845 | -86.238
1.2848 | -86.2334
1.2851 | -86.2306
1.2854 | -86.2301
1.2857 | -86.2274
1.286 | -86.2229
1.2863 | -86.2189
1.2866 | -86.2136
1.2869 | -86.2069
1.2873 | -86.2006
1.2876 | -86.1942
1.2879 | -86.1851
1.2882 | -86.1806
1.2885 | -86.1744
1.2888 | -86.1704
1.2891 | -86.17
1.2894 | -86.169
1.2897 | -86.1653
1.29 | -86.1611
1.2903 | -86.1562
1.2906 | -86.1518
1.291 | -86.148
1.2913 | -86.1444
1.2916 | -86.1448
1.2919 | -86.1407
1.2922 | -86.1338
1.2925 | -86.128
1.2928 | -86.1281
1.2931 | -86.1286
1.2934 | -86.1286
1.2937 | -86.1272
1.294 | -86.1262
1.2944 | -86.1257
1.2947 | -86.1235
1.295 | -86.1199
1.2953 | -86.1156
1.2956 | -86.1166
1.2959 | -86.1209
1.2962 | -86.1227
1.2965 | -86.1213
1.2968 | -86.1172
1.2971 | -86.1118
1.2974 | -86.1103
1.2978 | -86.1091
1.2981 | -86.1083
1.2984 | -86.1072
1.2987 | -86.1062
1.299 | -86.1064
1.2993 | -86.1055
1.2996 | -86.1068
1.2999 | -86.1069
1.3002 | -86.1054
1.3005 | -86.1121
1.3008 | -86.1145
1.3011 | -86.1174
1.3015 | -86.1252
1.3018 | -86.135
1.3021 | -86.1379
1.3024 | -86.137
1.3027 | -86.1385
1.303 | -86.1477
1.3033 | -86.1497
1.3036 | -86.1469
1.3039 | -86.1488
1.3042 | -86.1535
1.3045 | -86.1541
1.3049 | -86.1553
1.3052 | -86.1567
1.3055 | -86.1563
1.3058 | -86.1587
1.3061 | -86.1628
1.3064 | -86.1648
1.3067 | -86.1681
1.307 | -86.1717
1.3073 | -86.1762
1.3076 | -86.181
1.3079 | -86.1856
1.3082 | -86.1888
1.3086 | -86.1935
1.3089 | -86.199
1.3092 | -86.2036
1.3095 | -86.2109
1.3098 | -86.2168
1.3101 | -86.2228
1.3104 | -86.2259
1.3107 | -86.2313
1.311 | -86.2367
1.3113 | -86.2417
1.3116 | -86.2476
1.312 | -86.2551
1.3123 | -86.2637
1.3126 | -86.2712
1.3129 | -86.2784
1.3132 | -86.2853
1.3135 | -86.2934
1.3138 | -86.3012
1.3141 | -86.3083
1.3144 | -86.3173
1.3147 | -86.3254
1.315 | -86.3348
1.3154 | -86.3459
1.3157 | -86.3569
1.316 | -86.3669
1.3163 | -86.3755
1.3166 | -86.3836
1.3169 | -86.3918
1.3172 | -86.3998
1.3175 | -86.4085
1.3178 | -86.4192
1.3181 | -86.4289
1.3184 | -86.4437
1.3187 | -86.4557
1.3191 | -86.466
1.3194 | -86.476
1.3197 | -86.4876
1.32 | -86.4997
1.3203 | -86.5096
1.3206 | -86.5245
1.3209 | -86.5349
1.3212 | -86.5444
1.3215 | -86.5549
1.3218 | -86.5614
1.3221 | -86.5721
1.3225 | -86.5826
1.3228 | -86.5933
1.3231 | -86.6054
1.3234 | -86.6166
1.3237 | -86.6277
1.324 | -86.6411
1.3243 | -86.6539
1.3246 | -86.6648
1.3249 | -86.6758
1.3252 | -86.687
1.3255 | -86.7025
1.3258 | -86.7163
1.3262 | -86.7287
1.3265 | -86.7417
1.3268 | -86.7571
1.3271 | -86.7687
1.3274 | -86.7813
1.3277 | -86.7924
1.328 | -86.8025
1.3283 | -86.8111
1.3286 | -86.821
1.3289 | -86.8308
1.3292 | -86.8462
1.3296 | -86.8591
1.3299 | -86.8681
1.3302 | -86.8766
1.3305 | -86.8883
1.3308 | -86.898
1.3311 | -86.9098
1.3314 | -86.9206
1.3317 | -86.932
1.332 | -86.9436
1.3323 | -86.9547
1.3326 | -86.9677
1.333 | -86.9808
1.3333 | -86.9931
1.3336 | -87.0126
1.3339 | -87.0304
1.3342 | -87.0404
1.3345 | -87.0498
1.3348 | -87.0649
1.3351 | -87.0787
1.3354 | -87.0899
1.3357 | -87.1034
1.336 | -87.1117
1.3363 | -87.1214
1.3367 | -87.1305
1.337 | -87.1383
1.3373 | -87.1492
1.3376 | -87.1589
1.3379 | -87.1714
1.3382 | -87.1858
1.3385 | -87.1977
1.3388 | -87.2043
1.3391 | -87.2106
1.3394 | -87.2208
1.3397 | -87.2301
1.3401 | -87.2389
1.3404 | -87.2472
1.3407 | -87.2562
1.341 | -87.2655
1.3413 | -87.2784
1.3416 | -87.2899
1.3419 | -87.2944
1.3422 | -87.3
1.3425 | -87.3082
1.3428 | -87.3167
1.3431 | -87.3254
1.3434 | -87.3346
1.3438 | -87.3529
1.3441 | -87.3721
1.3444 | -87.382
1.3447 | -87.3881
1.345 | -87.3934
1.3453 | -87.4016
1.3456 | -87.409
1.3459 | -87.4164
1.3462 | -87.4256
1.3465 | -87.4364
1.3468 | -87.4469
1.3472 | -87.4522
1.3475 | -87.4619
1.3478 | -87.471
1.3481 | -87.4774
1.3484 | -87.4811
1.3487 | -87.4867
1.349 | -87.4934
1.3493 | -87.5038
1.3496 | -87.514
1.3499 | -87.5206
1.3502 | -87.5317
1.3505 | -87.5366
1.3509 | -87.5426
1.3512 | -87.5487
1.3515 | -87.5549
1.3518 | -87.5599
1.3521 | -87.5619
1.3524 | -87.57
1.3527 | -87.578
1.353 | -87.585
1.3533 | -87.5914
1.3536 | -87.5987
1.3539 | -87.5999
1.3543 | -87.6126
1.3546 | -87.6134
1.3549 | -87.6182
1.3552 | -87.6241
1.3555 | -87.6296
1.3558 | -87.6319
1.3561 | -87.6365
1.3564 | -87.6388
1.3567 | -87.6393
1.357 | -87.6425
1.3573 | -87.6419
1.3577 | -87.646
1.358 | -87.6498
1.3583 | -87.6529
1.3586 | -87.6556
1.3589 | -87.6617
1.3592 | -87.6659
1.3595 | -87.6676
1.3598 | -87.6663
1.3601 | -87.6705
1.3604 | -87.6763
1.3607 | -87.6802
1.361 | -87.6835
1.3614 | -87.6876
1.3617 | -87.69
1.362 | -87.6908
1.3623 | -87.6952
1.3626 | -87.7008
1.3629 | -87.7044
1.3632 | -87.708
1.3635 | -87.7052
1.3638 | -87.7062
1.3641 | -87.7076
1.3644 | -87.7111
1.3648 | -87.7141
1.3651 | -87.7152
1.3654 | -87.7143
1.3657 | -87.7153
1.366 | -87.7167
1.3663 | -87.7156
1.3666 | -87.721
1.3669 | -87.7317
1.3672 | -87.7381
1.3675 | -87.74
1.3678 | -87.7433
1.3681 | -87.7451
1.3685 | -87.7473
1.3688 | -87.7514
1.3691 | -87.7525
1.3694 | -87.748
1.3697 | -87.7478
1.37 | -87.75
1.3703 | -87.7509
1.3706 | -87.7515
1.3709 | -87.7513
1.3712 | -87.7504
1.3715 | -87.7504
1.3719 | -87.7434
1.3722 | -87.7413
1.3725 | -87.7414
1.3728 | -87.741
1.3731 | -87.7383
1.3734 | -87.7379
1.3737 | -87.7378
1.374 | -87.74
1.3743 | -87.7413
1.3746 | -87.7346
1.3749 | -87.7304
1.3753 | -87.731
1.3756 | -87.7253
1.3759 | -87.7242
1.3762 | -87.724
1.3765 | -87.7248
1.3768 | -87.7257
1.3771 | -87.7243
1.3774 | -87.7217
1.3777 | -87.7195
1.378 | -87.7167
1.3783 | -87.7138
1.3786 | -87.7124
1.379 | -87.714
1.3793 | -87.7136
1.3796 | -87.7115
1.3799 | -87.714
1.3802 | -87.7143
1.3805 | -87.7135
1.3808 | -87.7078
1.3811 | -87.7124
1.3814 | -87.7111
1.3817 | -87.7113
1.382 | -87.7115
1.3824 | -87.71
1.3827 | -87.7087
1.383 | -87.7067
1.3833 | -87.7023
1.3836 | -87.7005
1.3839 | -87.7002
1.3842 | -87.6987
1.3845 | -87.697
1.3848 | -87.6941
1.3851 | -87.6913
1.3854 | -87.6909
1.3858 | -87.6949
1.3861 | -87.699
1.3864 | -87.6972
1.3867 | -87.6957
1.387 | -87.6946
1.3873 | -87.693
1.3876 | -87.6911
1.3879 | -87.6897
1.3882 | -87.6884
1.3885 | -87.6888
1.3888 | -87.6903
1.3891 | -87.6888
1.3895 | -87.6891
1.3898 | -87.688
1.3901 | -87.6861
1.3904 | -87.685
1.3907 | -87.6857
1.391 | -87.6851
1.3913 | -87.6805
1.3916 | -87.6769
1.3919 | -87.6752
1.3922 | -87.6703
1.3925 | -87.6666
1.3929 | -87.6608
1.3932 | -87.6555
1.3935 | -87.6533
1.3938 | -87.6532
1.3941 | -87.6547
1.3944 | -87.6529
1.3947 | -87.6515
1.395 | -87.6465
1.3953 | -87.6432
1.3956 | -87.6399
1.3959 | -87.6367
1.3962 | -87.6296
1.3966 | -87.6194
1.3969 | -87.6162
1.3972 | -87.6165
1.3975 | -87.6192
1.3978 | -87.6206
1.3981 | -87.6214
1.3984 | -87.6209
1.3987 | -87.6188
1.399 | -87.6178
1.3993 | -87.6144
1.3996 | -87.6113
1.4 | -87.6114
1.4003 | -87.6074
1.4006 | -87.6043
1.4009 | -87.6022
1.4012 | -87.5956
1.4015 | -87.5918
1.4018 | -87.5891
1.4021 | -87.5869
1.4024 | -87.5847
1.4027 | -87.5852
1.403 | -87.5859
1.4034 | -87.5848
1.4037 | -87.582
1.404 | -87.5788
1.4043 | -87.5722
1.4046 | -87.5664
1.4049 | -87.5569
1.4052 | -87.5504
1.4055 | -87.5478
1.4058 | -87.544
1.4061 | -87.5405
1.4064 | -87.5328
1.4067 | -87.5264
1.4071 | -87.5214
1.4074 | -87.5134
1.4077 | -87.5057
1.408 | -87.5012
1.4083 | -87.4965
1.4086 | -87.4908
1.4089 | -87.4851
1.4092 | -87.4812
1.4095 | -87.4765
1.4098 | -87.4707
1.4101 | -87.4654
1.4105 | -87.4599
1.4108 | -87.4569
1.4111 | -87.4539
1.4114 | -87.4493
1.4117 | -87.4462
1.412 | -87.4433
1.4123 | -87.4408
1.4126 | -87.438
1.4129 | -87.435
1.4132 | -87.4313
1.4135 | -87.4267
1.4138 | -87.4225
1.4142 | -87.4182
1.4145 | -87.4147
1.4148 | -87.4117
1.4151 | -87.4054
1.4154 | -87.398
1.4157 | -87.3986
1.416 | -87.394
1.4163 | -87.3922
1.4166 | -87.3892
1.4169 | -87.3835
1.4172 | -87.3793
1.4176 | -87.3787
1.4179 | -87.3748
1.4182 | -87.371
1.4185 | -87.3672
1.4188 | -87.3624
1.4191 | -87.3587
1.4194 | -87.3555
1.4197 | -87.3503
1.42 | -87.3465
1.4203 | -87.3424
1.4206 | -87.3378
1.4209 | -87.3333
1.4213 | -87.3301
1.4216 | -87.3274
1.4219 | -87.3234
1.4222 | -87.3154
1.4225 | -87.3088
1.4228 | -87.3062
1.4231 | -87.3022
1.4234 | -87.2974
1.4237 | -87.2944
1.424 | -87.2939
1.4243 | -87.293
1.4247 | -87.29
1.425 | -87.2887
1.4253 | -87.2875
1.4256 | -87.2835
1.4259 | -87.2804
1.4262 | -87.2791
1.4265 | -87.2798
1.4268 | -87.2782
1.4271 | -87.2749
1.4274 | -87.2716
1.4277 | -87.2691
1.4281 | -87.2651
1.4284 | -87.2601
1.4287 | -87.2544
1.429 | -87.2422
1.4293 | -87.2329
1.4296 | -87.2287
1.4299 | -87.224
1.4302 | -87.219
1.4305 | -87.2171
1.4308 | -87.214
1.4311 | -87.2098
1.4314 | -87.2107
1.4318 | -87.2091
1.4321 | -87.2053
1.4324 | -87.2004
1.4327 | -87.1962
1.433 | -87.1932
1.4333 | -87.1908
1.4336 | -87.1897
1.4339 | -87.1862
1.4342 | -87.1826
1.4345 | -87.1762
1.4348 | -87.1724
1.4352 | -87.1685
1.4355 | -87.1629
1.4358 | -87.1606
1.4361 | -87.1579
1.4364 | -87.1532
1.4367 | -87.1488
1.437 | -87.1432
1.4373 | -87.1397
1.4376 | -87.1367
1.4379 | -87.1291
1.4382 | -87.1211
1.4385 | -87.1136
1.4389 | -87.1106
1.4392 | -87.1086
1.4395 | -87.1071
1.4398 | -87.1025
1.4401 | -87.0969
1.4404 | -87.0899
1.4407 | -87.0855
1.441 | -87.0823
1.4413 | -87.0782
1.4416 | -87.0727
1.4419 | -87.0629
1.4423 | -87.0531
1.4426 | -87.0442
1.4429 | -87.0351
1.4432 | -87.0317
1.4435 | -87.0281
1.4438 | -87.0225
1.4441 | -87.0173
1.4444 | -87.0145
1.4447 | -87.0145
1.445 | -87.0134
1.4453 | -87.0078
1.4457 | -87.002
1.446 | -86.9955
1.4463 | -86.9825
1.4466 | -86.9761
1.4469 | -86.9729
1.4472 | -86.969
1.4475 | -86.967
1.4478 | -86.9633
1.4481 | -86.9584
1.4484 | -86.9536
1.4487 | -86.9501
1.449 | -86.9443
1.4494 | -86.9394
1.4497 | -86.9355
1.45 | -86.9275
1.4503 | -86.9192
1.4506 | -86.9132
1.4509 | -86.9075
1.4512 | -86.9015
1.4515 | -86.899
1.4518 | -86.8952
1.4521 | -86.8895
1.4524 | -86.8813
1.4528 | -86.876
1.4531 | -86.8762
1.4534 | -86.8705
1.4537 | -86.8649
1.454 | -86.8605
1.4543 | -86.8525
1.4546 | -86.8437
1.4549 | -86.8368
1.4552 | -86.8294
1.4555 | -86.8272
1.4558 | -86.8203
1.4561 | -86.8132
1.4565 | -86.8048
1.4568 | -86.7972
1.4571 | -86.7909
1.4574 | -86.7829
1.4577 | -86.7736
1.458 | -86.771
1.4583 | -86.7658
1.4586 | -86.7568
1.4589 | -86.7505
1.4592 | -86.7438
1.4595 | -86.7349
1.4599 | -86.724
1.4602 | -86.7163
1.4605 | -86.7086
1.4608 | -86.7014
1.4611 | -86.6946
1.4614 | -86.6869
1.4617 | -86.6794
1.462 | -86.6742
1.4623 | -86.6714
1.4626 | -86.6647
1.4629 | -86.6576
1.4632 | -86.6507
1.4636 | -86.6438
1.4639 | -86.6367
1.4642 | -86.631
1.4645 | -86.6229
1.4648 | -86.6136
1.4651 | -86.6084
1.4654 | -86.5995
1.4657 | -86.5943
1.466 | -86.5878
1.4663 | -86.58
1.4666 | -86.5721
1.467 | -86.5652
1.4673 | -86.5576
1.4676 | -86.5499
1.4679 | -86.5419
1.4682 | -86.5379
1.4685 | -86.5314
1.4688 | -86.524
1.4691 | -86.5169
1.4694 | -86.509
1.4697 | -86.5023
1.47 | -86.4962
1.4704 | -86.4906
1.4707 | -86.4838
1.471 | -86.4791
1.4713 | -86.4756
1.4716 | -86.4682
1.4719 | -86.4578
1.4722 | -86.4512
1.4725 | -86.4468
1.4728 | -86.4411
1.4731 | -86.4356
1.4734 | -86.4296
1.4737 | -86.4214
1.4741 | -86.4179
1.4744 | -86.4125
1.4747 | -86.4053
1.475 | -86.3995
1.4753 | -86.3902
1.4756 | -86.3834
1.4759 | -86.3768
1.4762 | -86.366
1.4765 | -86.3535
1.4768 | -86.346
1.4771 | -86.3414
1.4775 | -86.3365
1.4778 | -86.3309
1.4781 | -86.3257
1.4784 | -86.3197
1.4787 | -86.3167
1.479 | -86.3105
1.4793 | -86.3037
1.4796 | -86.2969
1.4799 | -86.2919
1.4802 | -86.2862
1.4805 | -86.2822
1.4808 | -86.2783
1.4812 | -86.2724
1.4815 | -86.2665
1.4818 | -86.262
1.4821 | -86.2571
1.4824 | -86.2513
1.4827 | -86.2469
1.483 | -86.2429
1.4833 | -86.2367
1.4836 | -86.2296
1.4839 | -86.2242
1.4842 | -86.2198
1.4846 | -86.2154
1.4849 | -86.2119
1.4852 | -86.2087
1.4855 | -86.2053
1.4858 | -86.2022
1.4861 | -86.1991
1.4864 | -86.1962
1.4867 | -86.1927
1.487 | -86.1896
1.4873 | -86.1869
1.4876 | -86.1838
1.488 | -86.1792
1.4883 | -86.1766
1.4886 | -86.1737
1.4889 | -86.1717
1.4892 | -86.1708
1.4895 | -86.17
1.4898 | -86.1671
1.4901 | -86.1665
1.4904 | -86.1619
1.4907 | -86.1605
1.491 | -86.1588
1.4913 | -86.158
1.4917 | -86.1592
1.492 | -86.1613
1.4923 | -86.1626
1.4926 | -86.162
1.4929 | -86.1628
1.4932 | -86.1651
1.4935 | -86.1682
1.4938 | -86.1717
1.4941 | -86.1725
1.4944 | -86.1736
1.4947 | -86.1743
1.4951 | -86.1745
1.4954 | -86.1755
1.4957 | -86.1785
1.496 | -86.1818
1.4963 | -86.1845
1.4966 | -86.1864
1.4969 | -86.1892
1.4972 | -86.1948
1.4975 | -86.2036
1.4978 | -86.2079
1.4981 | -86.2104
1.4985 | -86.21
1.4988 | -86.2123
1.4991 | -86.214
1.4994 | -86.2165
1.4997 | -86.2192
1.5 | -86.221
1.5003 | -86.2211
1.5006 | -86.2244
1.5009 | -86.2289
1.5012 | -86.2326
1.5015 | -86.2371
1.5018 | -86.2414
1.5022 | -86.2452
1.5025 | -86.2492
1.5028 | -86.2515
1.5031 | -86.2546
1.5034 | -86.2571
1.5037 | -86.2612
1.504 | -86.266
1.5043 | -86.2686
1.5046 | -86.2735
1.5049 | -86.28
1.5052 | -86.2868
1.5056 | -86.2964
1.5059 | -86.3038
1.5062 | -86.3105
1.5065 | -86.32
1.5068 | -86.3291
1.5071 | -86.3342
1.5074 | -86.3372
1.5077 | -86.3406
1.508 | -86.3494
1.5083 | -86.3573
1.5086 | -86.363
1.5089 | -86.3697
1.5093 | -86.3786
1.5096 | -86.3869
1.5099 | -86.3935
1.5102 | -86.4005
1.5105 | -86.4075
1.5108 | -86.415
1.5111 | -86.4219
1.5114 | -86.4268
1.5117 | -86.4277
1.512 | -86.4334
1.5123 | -86.441
1.5127 | -86.447
1.513 | -86.4519
1.5133 | -86.4584
1.5136 | -86.4642
1.5139 | -86.4691
1.5142 | -86.4723
1.5145 | -86.4776
1.5148 | -86.4843
1.5151 | -86.4892
1.5154 | -86.4954
1.5157 | -86.5025
1.516 | -86.5092
1.5164 | -86.5147
1.5167 | -86.5174
1.517 | -86.5198
1.5173 | -86.5278
1.5176 | -86.5357
1.5179 | -86.5413
1.5182 | -86.5475
1.5185 | -86.5524
1.5188 | -86.5588
1.5191 | -86.5645
1.5194 | -86.5728
1.5198 | -86.5802
1.5201 | -86.5877
1.5204 | -86.5944
1.5207 | -86.5999
1.521 | -86.6061
1.5213 | -86.6127
1.5216 | -86.6164
1.5219 | -86.6215
1.5222 | -86.6291
1.5225 | -86.636
1.5228 | -86.642
1.5232 | -86.6469
1.5235 | -86.6529
1.5238 | -86.6578
1.5241 | -86.6631
1.5244 | -86.6681
1.5247 | -86.6743
1.525 | -86.6833
1.5253 | -86.6919
1.5256 | -86.6986
1.5259 | -86.7062
1.5262 | -86.713
1.5265 | -86.7187
1.5269 | -86.7242
1.5272 | -86.7302
1.5275 | -86.7383
1.5278 | -86.7431
1.5281 | -86.748
1.5284 | -86.7532
1.5287 | -86.7589
1.529 | -86.7643
1.5293 | -86.7698
1.5296 | -86.7758
1.5299 | -86.779
1.5303 | -86.7796
1.5306 | -86.7848
1.5309 | -86.7883
1.5312 | -86.7906
1.5315 | -86.7911
1.5318 | -86.7969
1.5321 | -86.8004
1.5324 | -86.8129
1.5327 | -86.8228
1.533 | -86.8279
1.5333 | -86.8332
1.5336 | -86.8367
1.534 | -86.8382
1.5343 | -86.839
1.5346 | -86.8417
1.5349 | -86.8445
1.5352 | -86.8483
1.5355 | -86.8519
1.5358 | -86.8565
1.5361 | -86.8616
1.5364 | -86.8647
1.5367 | -86.8677
1.537 | -86.8705
1.5374 | -86.8734
1.5377 | -86.8774
1.538 | -86.8816
1.5383 | -86.8873
1.5386 | -86.8898
1.5389 | -86.8931
1.5392 | -86.8976
1.5395 | -86.9015
1.5398 | -86.907
1.5401 | -86.9105
1.5404 | -86.9144
1.5408 | -86.9189
1.5411 | -86.9203
1.5414 | -86.9239
1.5417 | -86.9265
1.542 | -86.9331
1.5423 | -86.9401
1.5426 | -86.9446
1.5429 | -86.9483
1.5432 | -86.9523
1.5435 | -86.9501
1.5438 | -86.9503
1.5441 | -86.9538
1.5445 | -86.9538
1.5448 | -86.9476
1.5451 | -86.9441
1.5454 | -86.9446
1.5457 | -86.949
1.546 | -86.957
1.5463 | -86.9633
1.5466 | -86.9673
1.5469 | -86.9691
1.5472 | -86.9655
1.5475 | -86.9644
1.5479 | -86.9637
1.5482 | -86.9647
1.5485 | -86.9647
1.5488 | -86.9642
1.5491 | -86.966
1.5494 | -86.9653
1.5497 | -86.965
1.55 | -86.9631
1.5503 | -86.9617
1.5506 | -86.9606
1.5509 | -86.9667
1.5512 | -86.9675
1.5516 | -86.9666
1.5519 | -86.9661
1.5522 | -86.9641
1.5525 | -86.9602
1.5528 | -86.9588
1.5531 | -86.9565
1.5534 | -86.9536
1.5537 | -86.9514
1.554 | -86.9475
1.5543 | -86.9463
1.5546 | -86.9477
1.555 | -86.9472
1.5553 | -86.9474
1.5556 | -86.9458
1.5559 | -86.9456
1.5562 | -86.9404
1.5565 | -86.9372
1.5568 | -86.9357
1.5571 | -86.9356
1.5574 | -86.9375
1.5577 | -86.9389
1.558 | -86.9374
1.5584 | -86.9383
1.5587 | -86.9368
1.559 | -86.9363
1.5593 | -86.9347
1.5596 | -86.9352
1.5599 | -86.9375
1.5602 | -86.9397
1.5605 | -86.9399
1.5608 | -86.9389
1.5611 | -86.9367
1.5614 | -86.9338
1.5617 | -86.9344
1.5621 | -86.9297
1.5624 | -86.9251
1.5627 | -86.9251
1.563 | -86.927
1.5633 | -86.9244
1.5636 | -86.9221
1.5639 | -86.9226
1.5642 | -86.9209
1.5645 | -86.9192
1.5648 | -86.9189
1.5651 | -86.9191
1.5655 | -86.9171
1.5658 | -86.9141
1.5661 | -86.9102
1.5664 | -86.908
1.5667 | -86.9061
1.567 | -86.9022
1.5673 | -86.9001
1.5676 | -86.8988
1.5679 | -86.8916
1.5682 | -86.8819
1.5685 | -86.8788
1.5688 | -86.8768
1.5692 | -86.8735
1.5695 | -86.8691
1.5698 | -86.8647
1.5701 | -86.8617
1.5704 | -86.8605
1.5707 | -86.8581
1.571 | -86.8546
1.5713 | -86.8534
1.5716 | -86.8504
1.5719 | -86.8487
1.5722 | -86.8452
1.5726 | -86.8428
1.5729 | -86.8401
1.5732 | -86.8375
1.5735 | -86.8351
1.5738 | -86.8337
1.5741 | -86.8314
1.5744 | -86.8281
1.5747 | -86.8218
1.575 | -86.8186
1.5753 | -86.8171
1.5756 | -86.8144
1.576 | -86.8075
1.5763 | -86.804
1.5766 | -86.8
1.5769 | -86.7964
1.5772 | -86.7919
1.5775 | -86.7861
1.5778 | -86.7832
1.5781 | -86.7833
1.5784 | -86.7842
1.5787 | -86.7812
1.579 | -86.7777
1.5793 | -86.7746
1.5797 | -86.7708
1.58 | -86.7677
1.5803 | -86.7656
1.5806 | -86.7637
1.5809 | -86.7632
1.5812 | -86.7611
1.5815 | -86.7591
1.5818 | -86.7587
1.5821 | -86.7502
1.5824 | -86.7411
1.5827 | -86.7408
1.5831 | -86.7427
1.5834 | -86.7406
1.5837 | -86.7369
1.584 | -86.7337
1.5843 | -86.7302
1.5846 | -86.727
1.5849 | -86.7226
1.5852 | -86.7192
1.5855 | -86.7174
1.5858 | -86.7093
1.5861 | -86.7062
1.5864 | -86.7056
1.5868 | -86.702
1.5871 | -86.6954
1.5874 | -86.6932
1.5877 | -86.6892
1.588 | -86.6795
1.5883 | -86.6738
1.5886 | -86.6692
1.5889 | -86.6643
1.5892 | -86.6601
1.5895 | -86.6571
1.5898 | -86.6534
1.5902 | -86.65
1.5905 | -86.6499
1.5908 | -86.6501
1.5911 | -86.6476
1.5914 | -86.6422
1.5917 | -86.6377
1.592 | -86.6335
1.5923 | -86.6273
1.5926 | -86.6217
1.5929 | -86.6183
1.5932 | -86.6146
1.5936 | -86.6115
1.5939 | -86.6088
1.5942 | -86.6048
1.5945 | -86.6037
1.5948 | -86.6003
1.5951 | -86.5984
1.5954 | -86.5986
1.5957 | -86.5975
1.596 | -86.5947
1.5963 | -86.5915
1.5966 | -86.5886
1.5969 | -86.585
1.5973 | -86.5817
1.5976 | -86.5808
1.5979 | -86.5779
1.5982 | -86.5778
1.5985 | -86.5727
1.5988 | -86.5686
1.5991 | -86.5638
1.5994 | -86.5542
1.5997 | -86.5451
1.6 | -86.5342
1.6003 | -86.5244
1.6007 | -86.5196
1.601 | -86.5163
1.6013 | -86.512
1.6016 | -86.5076
1.6019 | -86.5046
1.6022 | -86.501
1.6025 | -86.4982
1.6028 | -86.4954
1.6031 | -86.4932
1.6034 | -86.4901
1.6037 | -86.4873
1.604 | -86.4845
1.6044 | -86.4818
1.6047 | -86.4812
1.605 | -86.4808
1.6053 | -86.4782
1.6056 | -86.4762
1.6059 | -86.4727
1.6062 | -86.469
1.6065 | -86.4669
1.6068 | -86.467
1.6071 | -86.4649
1.6074 | -86.4624
1.6078 | -86.4603
1.6081 | -86.4578
1.6084 | -86.4541
1.6087 | -86.4542
1.609 | -86.4546
1.6093 | -86.4527
1.6096 | -86.4455
1.6099 | -86.4344
1.6102 | -86.4292
1.6105 | -86.4261
1.6108 | -86.4254
1.6112 | -86.423
1.6115 | -86.4213
1.6118 | -86.4168
1.6121 | -86.4106
1.6124 | -86.4086
1.6127 | -86.406
1.613 | -86.4054
1.6133 | -86.4039
1.6136 | -86.4004
1.6139 | -86.3986
1.6142 | -86.3964
1.6145 | -86.394
1.6149 | -86.3877
1.6152 | -86.3843
1.6155 | -86.3818
1.6158 | -86.3793
1.6161 | -86.3771
1.6164 | -86.377
1.6167 | -86.3763
1.617 | -86.3736
1.6173 | -86.3698
1.6176 | -86.3649
1.6179 | -86.3623
1.6183 | -86.3602
1.6186 | -86.3589
1.6189 | -86.3572
1.6192 | -86.3552
1.6195 | -86.3531
1.6198 | -86.351
1.6201 | -86.3487
1.6204 | -86.3455
1.6207 | -86.3481
1.621 | -86.3472
1.6213 | -86.3478
1.6216 | -86.3456
1.622 | -86.3431
1.6223 | -86.3395
1.6226 | -86.3343
1.6229 | -86.3334
1.6232 | -86.3316
1.6235 | -86.3328
1.6238 | -86.3303
1.6241 | -86.3283
1.6244 | -86.3241
1.6247 | -86.3221
1.625 | -86.3194
1.6254 | -86.3168
1.6257 | -86.3152
1.626 | -86.3148
1.6263 | -86.3164
1.6266 | -86.3169
1.6269 | -86.3155
1.6272 | -86.3147
1.6275 | -86.3112
1.6278 | -86.3079
1.6281 | -86.3053
1.6284 | -86.302
1.6287 | -86.3011
1.6291 | -86.2987
1.6294 | -86.2932
1.6297 | -86.2912
1.63 | -86.29
1.6303 | -86.2869
1.6306 | -86.2856
1.6309 | -86.284
1.6312 | -86.2823
1.6315 | -86.2829
1.6318 | -86.2817
1.6321 | -86.2795
1.6325 | -86.2775
1.6328 | -86.2771
1.6331 | -86.2761
1.6334 | -86.274
1.6337 | -86.2718
1.634 | -86.2718
1.6343 | -86.2707
1.6346 | -86.2701
1.6349 | -86.2687
1.6352 | -86.266
1.6355 | -86.2584
1.6359 | -86.2529
1.6362 | -86.251
1.6365 | -86.2494
1.6368 | -86.2482
1.6371 | -86.2467
1.6374 | -86.2459
1.6377 | -86.2442
1.638 | -86.2425
1.6383 | -86.2413
1.6386 | -86.2413
1.6389 | -86.2429
1.6392 | -86.2408
1.6396 | -86.2395
1.6399 | -86.2388
1.6402 | -86.237
1.6405 | -86.2366
1.6408 | -86.2392
1.6411 | -86.2369
1.6414 | -86.2292
1.6417 | -86.2263
1.642 | -86.229
1.6423 | -86.2239
1.6426 | -86.219
1.643 | -86.2173
1.6433 | -86.2154
1.6436 | -86.2141
1.6439 | -86.2127
1.6442 | -86.2118
1.6445 | -86.2092
1.6448 | -86.2097
1.6451 | -86.2117
1.6454 | -86.2104
1.6457 | -86.2092
1.646 | -86.2087
1.6463 | -86.2071
1.6467 | -86.2056
1.647 | -86.2033
1.6473 | -86.2029
1.6476 | -86.201
1.6479 | -86.1997
1.6482 | -86.1993
1.6485 | -86.199
1.6488 | -86.1984
1.6491 | -86.1975
1.6494 | -86.1939
1.6497 | -86.1943
1.6501 | -86.1902
1.6504 | -86.1879
1.6507 | -86.1866
1.651 | -86.1857
1.6513 | -86.1839
1.6516 | -86.1787
1.6519 | -86.1748
1.6522 | -86.176
1.6525 | -86.1702
1.6528 | -86.1663
1.6531 | -86.1652
1.6535 | -86.1625
1.6538 | -86.16
1.6541 | -86.1608
1.6544 | -86.1616
1.6547 | -86.1612
1.655 | -86.1612
1.6553 | -86.1619
1.6556 | -86.1614
1.6559 | -86.1653
1.6562 | -86.172
1.6565 | -86.1744
1.6568 | -86.1766
1.6572 | -86.1769
1.6575 | -86.1758
1.6578 | -86.1755
1.6581 | -86.1739
1.6584 | -86.1731
1.6587 | -86.1753
1.659 | -86.179
1.6593 | -86.1791
1.6596 | -86.1787
1.6599 | -86.1775
1.6602 | -86.1766
1.6606 | -86.1755
1.6609 | -86.1749
1.6612 | -86.1709
1.6615 | -86.1679
1.6618 | -86.1631
1.6621 | -86.1637
1.6624 | -86.1617
1.6627 | -86.1617
1.663 | -86.1616
1.6633 | -86.1606
1.6636 | -86.1593
1.664 | -86.1586
1.6643 | -86.1581
1.6646 | -86.1539
1.6649 | -86.1475
1.6652 | -86.1386
1.6655 | -86.136
1.6658 | -86.1352
1.6661 | -86.1345
1.6664 | -86.1334
1.6667 | -86.1314
1.667 | -86.1306
1.6673 | -86.1307
1.6677 | -86.1278
1.668 | -86.1275
1.6683 | -86.1265
1.6686 | -86.1267
1.6689 | -86.1263
1.6692 | -86.1232
1.6695 | -86.1186
1.6698 | -86.1159
1.6701 | -86.1146
1.6704 | -86.1121
1.6707 | -86.1096
1.6711 | -86.1084
1.6714 | -86.1075
1.6717 | -86.1054
1.672 | -86.1032
1.6723 | -86.1018
1.6726 | -86.1001
1.6729 | -86.0977
1.6732 | -86.0939
1.6735 | -86.0913
1.6738 | -86.0918
1.6741 | -86.0921
1.6744 | -86.0886
1.6748 | -86.0858
1.6751 | -86.0808
1.6754 | -86.0726
1.6757 | -86.0677
1.676 | -86.0659
1.6763 | -86.0608
1.6766 | -86.0563
1.6769 | -86.054
1.6772 | -86.0511
1.6775 | -86.0483
1.6778 | -86.0459
1.6782 | -86.044
1.6785 | -86.0432
1.6788 | -86.041
1.6791 | -86.0386
1.6794 | -86.0365
1.6797 | -86.0318
1.68 | -86.0299
1.6803 | -86.0271
1.6806 | -86.024
1.6809 | -86.0191
1.6812 | -86.0166
1.6815 | -86.017
1.6819 | -86.0142
1.6822 | -86.0092
1.6825 | -86.0068
1.6828 | -85.9984
1.6831 | -85.9901
1.6834 | -85.9847
1.6837 | -85.9771
1.684 | -85.9695
1.6843 | -85.9647
1.6846 | -85.9604
1.6849 | -85.9548
1.6853 | -85.9499
1.6856 | -85.9452
1.6859 | -85.9406
1.6862 | -85.9348
1.6865 | -85.9302
1.6868 | -85.9262
1.6871 | -85.9213
1.6874 | -85.9168
1.6877 | -85.9095
1.688 | -85.9024
1.6883 | -85.8991
1.6887 | -85.8935
1.689 | -85.8869
1.6893 | -85.882
1.6896 | -85.8796
1.6899 | -85.8727
1.6902 | -85.8667
1.6905 | -85.863
1.6908 | -85.8582
1.6911 | -85.8535
1.6914 | -85.8489
1.6917 | -85.8447
1.692 | -85.839
1.6924 | -85.8326
1.6927 | -85.8252
1.693 | -85.8182
1.6933 | -85.8117
1.6936 | -85.8052
1.6939 | -85.8001
1.6942 | -85.7941
1.6945 | -85.7853
1.6948 | -85.7783
1.6951 | -85.7729
1.6954 | -85.7678
1.6958 | -85.7629
1.6961 | -85.7575
1.6964 | -85.7539
1.6967 | -85.7467
1.697 | -85.7387
1.6973 | -85.7334
1.6976 | -85.7256
1.6979 | -85.7168
1.6982 | -85.7099
1.6985 | -85.7036
1.6988 | -85.6954
1.6991 | -85.6859
1.6995 | -85.6803
1.6998 | -85.6737
1.7001 | -85.664
1.7004 | -85.6596
1.7007 | -85.654
1.701 | -85.6477
1.7013 | -85.641
1.7016 | -85.6319
1.7019 | -85.6233
1.7022 | -85.6143
1.7025 | -85.6033
1.7029 | -85.5985
1.7032 | -85.5933
1.7035 | -85.5873
1.7038 | -85.581
1.7041 | -85.5731
1.7044 | -85.5661
1.7047 | -85.5589
1.705 | -85.5516
1.7053 | -85.5432
1.7056 | -85.5316
1.7059 | -85.5189
1.7063 | -85.5107
1.7066 | -85.5028
1.7069 | -85.495
1.7072 | -85.4894
1.7075 | -85.4835
1.7078 | -85.4768
1.7081 | -85.4701
1.7084 | -85.4631
1.7087 | -85.4559
1.709 | -85.4503
1.7093 | -85.4436
1.7096 | -85.4365
1.71 | -85.4302
1.7103 | -85.4218
1.7106 | -85.4169
1.7109 | -85.4107
1.7112 | -85.4047
1.7115 | -85.399
1.7118 | -85.3935
1.7121 | -85.3871
1.7124 | -85.3812
1.7127 | -85.3745
1.713 | -85.3648
1.7134 | -85.3569
1.7137 | -85.3508
1.714 | -85.3448
1.7143 | -85.3366
1.7146 | -85.3305
1.7149 | -85.3295
1.7152 | -85.3274
1.7155 | -85.3262
1.7158 | -85.3248
1.7161 | -85.3198
1.7164 | -85.3115
1.7167 | -85.3051
1.7171 | -85.2979
1.7174 | -85.2908
1.7177 | -85.2846
1.718 | -85.2782
1.7183 | -85.2715
1.7186 | -85.2647
1.7189 | -85.2586
1.7192 | -85.2559
1.7195 | -85.2521
1.7198 | -85.2477
1.7201 | -85.2423
1.7205 | -85.2362
1.7208 | -85.2292
1.7211 | -85.2239
1.7214 | -85.2188
1.7217 | -85.2139
1.722 | -85.2099
1.7223 | -85.2046
1.7226 | -85.1964
1.7229 | -85.1896
1.7232 | -85.1823
1.7235 | -85.1774
1.7239 | -85.1722
1.7242 | -85.1641
1.7245 | -85.1598
1.7248 | -85.1551
1.7251 | -85.1492
1.7254 | -85.1441
1.7257 | -85.141
1.726 | -85.1407
1.7263 | -85.1377
1.7266 | -85.1323
1.7269 | -85.1282
1.7272 | -85.1252
1.7276 | -85.1215
1.7279 | -85.1199
1.7282 | -85.1197
1.7285 | -85.1187
1.7288 | -85.1159
1.7291 | -85.1127
1.7294 | -85.1072
1.7297 | -85.1026
1.73 | -85.0984
1.7303 | -85.0953
1.7306 | -85.0856
1.731 | -85.0791
1.7313 | -85.0751
1.7316 | -85.0721
1.7319 | -85.0665
1.7322 | -85.0619
1.7325 | -85.0588
1.7328 | -85.0558
1.7331 | -85.0524
1.7334 | -85.0497
1.7337 | -85.0469
1.734 | -85.0454
1.7343 | -85.0452
1.7347 | -85.0435
1.735 | -85.0405
1.7353 | -85.0377
1.7356 | -85.0354
1.7359 | -85.0332
1.7362 | -85.0319
1.7365 | -85.0298
1.7368 | -85.027
1.7371 | -85.0285
1.7374 | -85.0284
1.7377 | -85.0282
1.7381 | -85.0271
1.7384 | -85.0249
1.7387 | -85.0236
1.739 | -85.0224
1.7393 | -85.0217
1.7396 | -85.0229
1.7399 | -85.0242
1.7402 | -85.0219
1.7405 | -85.0224
1.7408 | -85.021
1.7411 | -85.0208
1.7414 | -85.0214
1.7418 | -85.023
1.7421 | -85.0232
1.7424 | -85.0232
1.7427 | -85.0215
1.743 | -85.0184
1.7433 | -85.0176
1.7436 | -85.0179
1.7439 | -85.0182
1.7442 | -85.0161
1.7445 | -85.0164
1.7448 | -85.0183
1.7452 | -85.0204
1.7455 | -85.022
1.7458 | -85.0237
1.7461 | -85.0278
1.7464 | -85.0279
1.7467 | -85.0289
1.747 | -85.0318
1.7473 | -85.0324
1.7476 | -85.0329
1.7479 | -85.0338
1.7482 | -85.0349
1.7486 | -85.0359
1.7489 | -85.0359
1.7492 | -85.0367
1.7495 | -85.0382
1.7498 | -85.0384
1.7501 | -85.0365
1.7504 | -85.0351
1.7507 | -85.0327
1.751 | -85.0325
1.7513 | -85.0342
1.7516 | -85.0373
1.7519 | -85.0383
1.7523 | -85.0401
1.7526 | -85.0418
1.7529 | -85.0438
1.7532 | -85.0449
1.7535 | -85.0492
1.7538 | -85.0535
1.7541 | -85.0534
1.7544 | -85.055
1.7547 | -85.0595
1.755 | -85.0622
1.7553 | -85.0638
1.7557 | -85.0611
1.756 | -85.0619
1.7563 | -85.0653
1.7566 | -85.068
1.7569 | -85.0687
1.7572 | -85.0701
1.7575 | -85.0725
1.7578 | -85.0749
1.7581 | -85.077
1.7584 | -85.0791
1.7587 | -85.0813
1.7591 | -85.0851
1.7594 | -85.0879
1.7597 | -85.0919
1.76 | -85.1004
1.7603 | -85.1061
1.7606 | -85.1076
1.7609 | -85.1094
1.7612 | -85.1139
1.7615 | -85.1173
1.7618 | -85.12
1.7621 | -85.1219
1.7624 | -85.1248
1.7628 | -85.1276
1.7631 | -85.1329
1.7634 | -85.1359
1.7637 | -85.1412
1.764 | -85.1448
1.7643 | -85.1488
1.7646 | -85.1512
1.7649 | -85.1539
1.7652 | -85.1585
1.7655 | -85.1621
1.7658 | -85.1668
1.7662 | -85.1703
1.7665 | -85.1741
1.7668 | -85.1771
1.7671 | -85.1799
1.7674 | -85.1824
1.7677 | -85.1853
1.768 | -85.1893
1.7683 | -85.1943
1.7686 | -85.1974
1.7689 | -85.1941
1.7692 | -85.1961
1.7695 | -85.1991
1.7699 | -85.2028
1.7702 | -85.2064
1.7705 | -85.2111
1.7708 | -85.2173
1.7711 | -85.2212
1.7714 | -85.2219
1.7717 | -85.228
1.772 | -85.2339
1.7723 | -85.2373
1.7726 | -85.2426
1.7729 | -85.2477
1.7733 | -85.2526
1.7736 | -85.2566
1.7739 | -85.2598
1.7742 | -85.2636
1.7745 | -85.2692
1.7748 | -85.2737
1.7751 | -85.2774
1.7754 | -85.2802
1.7757 | -85.2837
1.776 | -85.2877
1.7763 | -85.2907
1.7767 | -85.2946
1.777 | -85.2983
1.7773 | -85.3019
1.7776 | -85.3035
1.7779 | -85.306
1.7782 | -85.3111
1.7785 | -85.3147
1.7788 | -85.3182
1.7791 | -85.3226
1.7794 | -85.3273
1.7797 | -85.3321
1.78 | -85.3372
1.7804 | -85.3391
1.7807 | -85.3359
1.781 | -85.3398
1.7813 | -85.3463
1.7816 | -85.3408
1.7819 | -85.3443
1.7822 | -85.3455
1.7825 | -85.351
1.7828 | -85.3558
1.7831 | -85.3591
1.7834 | -85.3628
1.7838 | -85.3649
1.7841 | -85.3671
1.7844 | -85.3706
1.7847 | -85.3783
1.785 | -85.385
1.7853 | -85.3928
1.7856 | -85.3955
1.7859 | -85.3983
1.7862 | -85.401
1.7865 | -85.404
1.7868 | -85.4076
1.7871 | -85.4101
1.7875 | -85.4124
1.7878 | -85.414
1.7881 | -85.4124
1.7884 | -85.4102
1.7887 | -85.4118
1.789 | -85.4143
1.7893 | -85.4154
1.7896 | -85.4183
1.7899 | -85.4207
1.7902 | -85.4228
1.7905 | -85.425
1.7909 | -85.4267
1.7912 | -85.4256
1.7915 | -85.4245
1.7918 | -85.4254
1.7921 | -85.4254
1.7924 | -85.4231
1.7927 | -85.424
1.793 | -85.4251
1.7933 | -85.4266
1.7936 | -85.4282
1.7939 | -85.428
1.7942 | -85.4287
1.7946 | -85.4271
1.7949 | -85.4271
1.7952 | -85.43
1.7955 | -85.43
1.7958 | -85.43
1.7961 | -85.4331
1.7964 | -85.4401
1.7967 | -85.4423
1.797 | -85.4419
1.7973 | -85.4445
1.7976 | -85.4461
1.798 | -85.446
1.7983 | -85.4459
1.7986 | -85.445
1.7989 | -85.4448
1.7992 | -85.4456
1.7995 | -85.4442
1.7998 | -85.4441
1.8001 | -85.444
1.8004 | -85.4448
1.8007 | -85.4448
1.801 | -85.4457
1.8014 | -85.4451
1.8017 | -85.444
1.802 | -85.4432
1.8023 | -85.4452
1.8026 | -85.4477
1.8029 | -85.4473
1.8032 | -85.4456
1.8035 | -85.4459
1.8038 | -85.4438
1.8041 | -85.4439
1.8044 | -85.4476
1.8047 | -85.4482
1.8051 | -85.4459
1.8054 | -85.4468
1.8057 | -85.4487
1.806 | -85.448
1.8063 | -85.4457
1.8066 | -85.4453
1.8069 | -85.445
1.8072 | -85.4426
1.8075 | -85.4394
1.8078 | -85.4389
1.8081 | -85.4369
1.8085 | -85.4348
1.8088 | -85.4312
1.8091 | -85.4277
1.8094 | -85.4252
1.8097 | -85.4254
1.81 | -85.4249
1.8103 | -85.4225
1.8106 | -85.4218
1.8109 | -85.4215
1.8112 | -85.4203
1.8115 | -85.4187
1.8118 | -85.4189
1.8122 | -85.4155
1.8125 | -85.4084
1.8128 | -85.4028
1.8131 | -85.3987
1.8134 | -85.3963
1.8137 | -85.394
1.814 | -85.3922
1.8143 | -85.3903
1.8146 | -85.3894
1.8149 | -85.3889
1.8152 | -85.3862
1.8156 | -85.382
1.8159 | -85.3799
1.8162 | -85.3774
1.8165 | -85.3755
1.8168 | -85.3758
1.8171 | -85.3736
1.8174 | -85.3702
1.8177 | -85.3659
1.818 | -85.3635
1.8183 | -85.3617
1.8186 | -85.3595
1.819 | -85.3589
1.8193 | -85.3615
1.8196 | -85.3598
1.8199 | -85.3521
1.8202 | -85.3475
1.8205 | -85.3438
1.8208 | -85.3404
1.8211 | -85.3381
1.8214 | -85.3367
1.8217 | -85.3368
1.822 | -85.3328
1.8223 | -85.3296
1.8227 | -85.3301
1.823 | -85.3288
1.8233 | -85.3255
1.8236 | -85.3225
1.8239 | -85.321
1.8242 | -85.3182
1.8245 | -85.3188
1.8248 | -85.3198
1.8251 | -85.3182
1.8254 | -85.3162
1.8257 | -85.3129
1.8261 | -85.3081
1.8264 | -85.3025
1.8267 | -85.2953
1.827 | -85.2887
1.8273 | -85.2869
1.8276 | -85.2838
1.8279 | -85.281
1.8282 | -85.2774
1.8285 | -85.2749
1.8288 | -85.2759
1.8291 | -85.2781
1.8294 | -85.2773
1.8298 | -85.2761
1.8301 | -85.2747
1.8304 | -85.2725
1.8307 | -85.2689
1.831 | -85.262
1.8313 | -85.2576
1.8316 | -85.2546
1.8319 | -85.2513
1.8322 | -85.2458
1.8325 | -85.2381
1.8328 | -85.2279
1.8332 | -85.2163
1.8335 | -85.2117
1.8338 | -85.2097
1.8341 | -85.2058
1.8344 | -85.2024
1.8347 | -85.1985
1.835 | -85.1964
1.8353 | -85.1926
1.8356 | -85.189
1.8359 | -85.1863
1.8362 | -85.1832
1.8366 | -85.1803
1.8369 | -85.1794
1.8372 | -85.1759
1.8375 | -85.173
1.8378 | -85.1703
1.8381 | -85.1628
1.8384 | -85.1581
1.8387 | -85.1543
1.839 | -85.1496
1.8393 | -85.1451
1.8396 | -85.1412
1.8399 | -85.1391
1.8403 | -85.1374
1.8406 | -85.133
1.8409 | -85.1282
1.8412 | -85.1239
1.8415 | -85.1211
1.8418 | -85.1155
1.8421 | -85.1128
1.8424 | -85.11
1.8427 | -85.1067
1.843 | -85.1038
1.8433 | -85.0996
1.8437 | -85.0958
1.844 | -85.0918
1.8443 | -85.0886
1.8446 | -85.0857
1.8449 | -85.0832
1.8452 | -85.0822
1.8455 | -85.0786
1.8458 | -85.0738
1.8461 | -85.0683
1.8464 | -85.065
1.8467 | -85.0626
1.847 | -85.0612
1.8474 | -85.06
1.8477 | -85.0593
1.848 | -85.0569
1.8483 | -85.0511
1.8486 | -85.0495
1.8489 | -85.0475
1.8492 | -85.0449
1.8495 | -85.0415
1.8498 | -85.0354
1.8501 | -85.0332
1.8504 | -85.0318
1.8508 | -85.0291
1.8511 | -85.0262
1.8514 | -85.0225
1.8517 | -85.0187
1.852 | -85.0137
1.8523 | -85.0091
1.8526 | -85.003
1.8529 | -85.0008
1.8532 | -84.9971
1.8535 | -84.9932
1.8538 | -84.9892
1.8541 | -84.9861
1.8545 | -84.9845
1.8548 | -84.9808
1.8551 | -84.9783
1.8554 | -84.9753
1.8557 | -84.9776
1.856 | -84.9772
1.8563 | -84.9742
1.8566 | -84.9735
1.8569 | -84.9713
1.8572 | -84.9677
1.8575 | -84.9631
1.8579 | -84.9601
1.8582 | -84.9576
1.8585 | -84.9545
1.8588 | -84.9516
1.8591 | -84.9475
1.8594 | -84.9439
1.8597 | -84.9407
1.86 | -84.9365
1.8603 | -84.9314
1.8606 | -84.928
1.8609 | -84.9252
1.8613 | -84.9198
1.8616 | -84.9157
1.8619 | -84.915
1.8622 | -84.9129
1.8625 | -84.9107
1.8628 | -84.9081
1.8631 | -84.9056
1.8634 | -84.9029
1.8637 | -84.9001
1.864 | -84.8961
1.8643 | -84.8876
1.8646 | -84.8812
1.865 | -84.8771
1.8653 | -84.8737
1.8656 | -84.8709
1.8659 | -84.8675
1.8662 | -84.862
1.8665 | -84.8563
1.8668 | -84.8502
1.8671 | -84.8471
1.8674 | -84.843
1.8677 | -84.8393
1.868 | -84.8367
1.8684 | -84.8346
1.8687 | -84.8312
1.869 | -84.8259
1.8693 | -84.8202
1.8696 | -84.8169
1.8699 | -84.8157
1.8702 | -84.8124
1.8705 | -84.8087
1.8708 | -84.8076
1.8711 | -84.804
1.8714 | -84.7963
1.8718 | -84.7908
1.8721 | -84.7869
1.8724 | -84.7834
1.8727 | -84.7811
1.873 | -84.7765
1.8733 | -84.7711
1.8736 | -84.767
1.8739 | -84.765
1.8742 | -84.7624
1.8745 | -84.7598
1.8748 | -84.7567
1.8751 | -84.7549
1.8755 | -84.7528
1.8758 | -84.7494
1.8761 | -84.747
1.8764 | -84.7455
1.8767 | -84.7443
1.877 | -84.7446
1.8773 | -84.7428
1.8776 | -84.7396
1.8779 | -84.7365
1.8782 | -84.7347
1.8785 | -84.7312
1.8789 | -84.7276
1.8792 | -84.7247
1.8795 | -84.7232
1.8798 | -84.7222
1.8801 | -84.7213
1.8804 | -84.7179
1.8807 | -84.7163
1.881 | -84.7133
1.8813 | -84.7093
1.8816 | -84.7064
1.8819 | -84.7037
1.8822 | -84.7023
1.8826 | -84.7029
1.8829 | -84.7015
1.8832 | -84.6988
1.8835 | -84.696
1.8838 | -84.6944
1.8841 | -84.693
1.8844 | -84.6904
1.8847 | -84.6868
1.885 | -84.6834
1.8853 | -84.682
1.8856 | -84.681
1.886 | -84.6788
1.8863 | -84.6772
1.8866 | -84.6751
1.8869 | -84.6716
1.8872 | -84.6678
1.8875 | -84.6656
1.8878 | -84.6671
1.8881 | -84.6655
1.8884 | -84.6653
1.8887 | -84.6633
1.889 | -84.6614
1.8894 | -84.6601
1.8897 | -84.6582
1.89 | -84.6562
1.8903 | -84.6511
1.8906 | -84.6481
1.8909 | -84.6462
1.8912 | -84.6437
1.8915 | -84.6398
1.8918 | -84.6362
1.8921 | -84.634
1.8924 | -84.633
1.8927 | -84.6312
1.8931 | -84.6294
1.8934 | -84.6277
1.8937 | -84.6256
1.894 | -84.6253
1.8943 | -84.6262
1.8946 | -84.6259
1.8949 | -84.6247
1.8952 | -84.6245
1.8955 | -84.623
1.8958 | -84.6212
1.8961 | -84.6229
1.8965 | -84.6276
1.8968 | -84.6292
1.8971 | -84.6289
1.8974 | -84.6314
1.8977 | -84.6306
1.898 | -84.6284
1.8983 | -84.6265
1.8986 | -84.6256
1.8989 | -84.6267
1.8992 | -84.6281
1.8995 | -84.6286
1.8998 | -84.6279
1.9002 | -84.627
1.9005 | -84.6268
1.9008 | -84.6266
1.9011 | -84.6259
1.9014 | -84.6244
1.9017 | -84.6223
1.902 | -84.6197
1.9023 | -84.6168
1.9026 | -84.6154
1.9029 | -84.6155
1.9032 | -84.6154
1.9036 | -84.6153
1.9039 | -84.6144
1.9042 | -84.6141
1.9045 | -84.6154
1.9048 | -84.6141
1.9051 | -84.6124
1.9054 | -84.6118
1.9057 | -84.6142
1.906 | -84.6159
1.9063 | -84.6187
1.9066 | -84.6193
1.9069 | -84.6199
1.9073 | -84.6198
1.9076 | -84.6205
1.9079 | -84.6213
1.9082 | -84.6225
1.9085 | -84.6204
1.9088 | -84.6201
1.9091 | -84.6267
1.9094 | -84.6299
1.9097 | -84.6299
1.91 | -84.6325
1.9103 | -84.6368
1.9107 | -84.637
1.911 | -84.6382
1.9113 | -84.6396
1.9116 | -84.6408
1.9119 | -84.6402
1.9122 | -84.6399
1.9125 | -84.6414
1.9128 | -84.6424
1.9131 | -84.6435
1.9134 | -84.6441
1.9137 | -84.6434
1.9141 | -84.6441
1.9144 | -84.6419
1.9147 | -84.6358
1.915 | -84.6343
1.9153 | -84.635
1.9156 | -84.6355
1.9159 | -84.6371
1.9162 | -84.6394
1.9165 | -84.6394
1.9168 | -84.6368
1.9171 | -84.6369
1.9174 | -84.6402
1.9178 | -84.6407
1.9181 | -84.6402
1.9184 | -84.6412
1.9187 | -84.642
1.919 | -84.644
1.9193 | -84.6454
1.9196 | -84.6464
1.9199 | -84.6474
1.9202 | -84.6484
1.9205 | -84.6508
1.9208 | -84.654
1.9212 | -84.6549
1.9215 | -84.656
1.9218 | -84.6604
1.9221 | -84.6615
1.9224 | -84.6618
1.9227 | -84.6604
1.923 | -84.6576
1.9233 | -84.658
1.9236 | -84.6597
1.9239 | -84.6614
1.9242 | -84.6626
1.9245 | -84.6642
1.9249 | -84.6647
1.9252 | -84.6663
1.9255 | -84.6693
1.9258 | -84.6726
1.9261 | -84.6736
1.9264 | -84.6748
1.9267 | -84.675
1.927 | -84.6754
1.9273 | -84.6769
1.9276 | -84.6776
1.9279 | -84.6778
1.9283 | -84.6794
1.9286 | -84.6804
1.9289 | -84.6792
1.9292 | -84.6782
1.9295 | -84.6799
1.9298 | -84.6787
1.9301 | -84.6769
1.9304 | -84.6774
1.9307 | -84.6764
1.931 | -84.6756
1.9313 | -84.6774
1.9317 | -84.6791
1.932 | -84.6797
1.9323 | -84.6798
1.9326 | -84.6802
1.9329 | -84.6818
1.9332 | -84.683
1.9335 | -84.6844
1.9338 | -84.6852
1.9341 | -84.6869
1.9344 | -84.6912
1.9347 | -84.6926
1.935 | -84.6929
1.9354 | -84.6941
1.9357 | -84.6954
1.936 | -84.6954
1.9363 | -84.6968
1.9366 | -84.6984
1.9369 | -84.6985
1.9372 | -84.6985
1.9375 | -84.6986
1.9378 | -84.6998
1.9381 | -84.6996
1.9384 | -84.6995
1.9388 | -84.7004
1.9391 | -84.7011
1.9394 | -84.7033
1.9397 | -84.7045
1.94 | -84.7045
1.9403 | -84.7047
1.9406 | -84.7033
1.9409 | -84.7015
1.9412 | -84.7006
1.9415 | -84.7009
1.9418 | -84.7018
1.9421 | -84.7023
1.9425 | -84.7022
1.9428 | -84.7022
1.9431 | -84.7036
1.9434 | -84.7073
1.9437 | -84.7104
1.944 | -84.7128
1.9443 | -84.7142
1.9446 | -84.7144
1.9449 | -84.7147
1.9452 | -84.7143
1.9455 | -84.7096
1.9459 | -84.7045
1.9462 | -84.704
1.9465 | -84.7063
1.9468 | -84.7064
1.9471 | -84.7064
1.9474 | -84.7065
1.9477 | -84.7052
1.948 | -84.7018
1.9483 | -84.7012
1.9486 | -84.703
1.9489 | -84.7024
1.9493 | -84.7
1.9496 | -84.6996
1.9499 | -84.6979
1.9502 | -84.6927
1.9505 | -84.6897
1.9508 | -84.6861
1.9511 | -84.685
1.9514 | -84.688
1.9517 | -84.6897
1.952 | -84.6914
1.9523 | -84.6921
1.9526 | -84.6894
1.953 | -84.6856
1.9533 | -84.6829
1.9536 | -84.6805
1.9539 | -84.6771
1.9542 | -84.6742
1.9545 | -84.671
1.9548 | -84.6691
1.9551 | -84.6674
1.9554 | -84.6638
1.9557 | -84.6613
1.956 | -84.6596
1.9564 | -84.6585
1.9567 | -84.6586
1.957 | -84.6605
1.9573 | -84.6616
1.9576 | -84.6606
1.9579 | -84.6626
1.9582 | -84.6638
1.9585 | -84.6657
1.9588 | -84.6662
1.9591 | -84.6648
1.9594 | -84.6636
1.9597 | -84.6594
1.9601 | -84.6564
1.9604 | -84.656
1.9607 | -84.6547
1.961 | -84.6508
1.9613 | -84.645
1.9616 | -84.6418
1.9619 | -84.6403
1.9622 | -84.6388
1.9625 | -84.6372
1.9628 | -84.6362
1.9631 | -84.6372
1.9635 | -84.6353
1.9638 | -84.6327
1.9641 | -84.6318
1.9644 | -84.6303
1.9647 | -84.6325
1.965 | -84.6369
1.9653 | -84.6424
1.9656 | -84.6423
1.9659 | -84.6389
1.9662 | -84.6373
1.9665 | -84.6356
1.9669 | -84.6344
1.9672 | -84.6334
1.9675 | -84.6306
1.9678 | -84.6269
1.9681 | -84.6252
1.9684 | -84.6214
1.9687 | -84.6124
1.969 | -84.6101
1.9693 | -84.6085
1.9696 | -84.6062
1.9699 | -84.605
1.9702 | -84.6035
1.9706 | -84.6022
1.9709 | -84.6054
1.9712 | -84.6056
1.9715 | -84.6027
1.9718 | -84.6015
1.9721 | -84.5989
1.9724 | -84.5924
1.9727 | -84.5878
1.973 | -84.585
1.9733 | -84.5828
1.9736 | -84.5814
1.974 | -84.5794
1.9743 | -84.5758
1.9746 | -84.5761
1.9749 | -84.573
1.9752 | -84.5693
1.9755 | -84.5659
1.9758 | -84.5641
1.9761 | -84.5627
1.9764 | -84.5606
1.9767 | -84.5599
1.977 | -84.5585
1.9773 | -84.5563
1.9777 | -84.5548
1.978 | -84.5527
1.9783 | -84.5504
1.9786 | -84.5474
1.9789 | -84.5459
1.9792 | -84.5447
1.9795 | -84.5442
1.9798 | -84.5416
1.9801 | -84.5406
1.9804 | -84.5398
1.9807 | -84.5382
1.9811 | -84.5366
1.9814 | -84.5348
1.9817 | -84.5354
1.982 | -84.5341
1.9823 | -84.5325
1.9826 | -84.53
1.9829 | -84.5296
1.9832 | -84.5268
1.9835 | -84.5188
1.9838 | -84.5156
1.9841 | -84.5128
1.9845 | -84.5116
1.9848 | -84.5105
1.9851 | -84.509
1.9854 | -84.5086
1.9857 | -84.5077
1.986 | -84.5073
1.9863 | -84.5054
1.9866 | -84.5045
1.9869 | -84.5026
1.9872 | -84.4995
1.9875 | -84.4963
1.9878 | -84.4948
1.9882 | -84.4932
1.9885 | -84.4916
1.9888 | -84.4895
1.9891 | -84.4867
1.9894 | -84.4861
1.9897 | -84.4853
1.99 | -84.4824
1.9903 | -84.4792
1.9906 | -84.4766
1.9909 | -84.4756
1.9912 | -84.4729
1.9916 | -84.472
1.9919 | -84.4702
1.9922 | -84.468
1.9925 | -84.4668
1.9928 | -84.4649
1.9931 | -84.4633
1.9934 | -84.4615
1.9937 | -84.4577
1.994 | -84.456
1.9943 | -84.4549
1.9946 | -84.4499
1.9949 | -84.4465
1.9953 | -84.4446
1.9956 | -84.4436
1.9959 | -84.4432
1.9962 | -84.4405
1.9965 | -84.4369
1.9968 | -84.4307
1.9971 | -84.428
1.9974 | -84.4263
1.9977 | -84.4251
1.998 | -84.4237
1.9983 | -84.4216
1.9987 | -84.4217
1.999 | -84.425
1.9993 | -84.4259
1.9996 | -84.4249
1.9999 | -84.4232
2.0002 | -84.4203
2.0005 | -84.4191
2.0008 | -84.4199
2.0011 | -84.4203
2.0014 | -84.4195
2.0017 | -84.4162
2.0021 | -84.4127
2.0024 | -84.4069
2.0027 | -84.4007
2.003 | -84.3961
2.0033 | -84.3951
2.0036 | -84.395
2.0039 | -84.3936
2.0042 | -84.3898
2.0045 | -84.3879
2.0048 | -84.3854
2.0051 | -84.3829
2.0054 | -84.3813
2.0058 | -84.3792
2.0061 | -84.3759
2.0064 | -84.3742
2.0067 | -84.3764
2.007 | -84.3765
2.0073 | -84.3752
2.0076 | -84.3733
2.0079 | -84.371
2.0082 | -84.3701
2.0085 | -84.3695
2.0088 | -84.3671
2.0092 | -84.3645
2.0095 | -84.3624
2.0098 | -84.3587
2.0101 | -84.357
2.0104 | -84.3553
2.0107 | -84.3533
2.011 | -84.3516
2.0113 | -84.3492
2.0116 | -84.3481
2.0119 | -84.3474
2.0122 | -84.3452
2.0125 | -84.3429
2.0129 | -84.3407
2.0132 | -84.3404
2.0135 | -84.3388
2.0138 | -84.3273
2.0141 | -84.3125
2.0144 | -84.3035
2.0147 | -84.2974
2.015 | -84.2929
2.0153 | -84.2886
2.0156 | -84.2843
2.0159 | -84.281
2.0163 | -84.2798
2.0166 | -84.2783
2.0169 | -84.2766
2.0172 | -84.2752
2.0175 | -84.273
2.0178 | -84.2703
2.0181 | -84.2676
2.0184 | -84.2651
2.0187 | -84.2629
2.019 | -84.2602
2.0193 | -84.2579
2.0196 | -84.2559
2.02 | -84.2537
2.0203 | -84.2517
2.0206 | -84.2514
2.0209 | -84.2502
2.0212 | -84.2462
2.0215 | -84.244
2.0218 | -84.242
2.0221 | -84.2388
2.0224 | -84.2367
2.0227 | -84.2322
2.023 | -84.2294
2.0234 | -84.2286
2.0237 | -84.227
2.024 | -84.223
2.0243 | -84.2175
2.0246 | -84.2159
2.0249 | -84.2138
2.0252 | -84.2107
2.0255 | -84.208
2.0258 | -84.2059
2.0261 | -84.1998
2.0264 | -84.1955
2.0268 | -84.1953
2.0271 | -84.192
2.0274 | -84.1878
2.0277 | -84.1862
2.028 | -84.1856
2.0283 | -84.1834
2.0286 | -84.1807
2.0289 | -84.1761
2.0292 | -84.1722
2.0295 | -84.1698
2.0298 | -84.1673
2.0301 | -84.1649
2.0305 | -84.1603
2.0308 | -84.1571
2.0311 | -84.1541
2.0314 | -84.1522
2.0317 | -84.1506
2.032 | -84.1477
2.0323 | -84.1453
2.0326 | -84.1436
2.0329 | -84.1408
2.0332 | -84.1379
2.0335 | -84.1362
2.0339 | -84.1345
2.0342 | -84.1313
2.0345 | -84.1275
2.0348 | -84.1248
2.0351 | -84.1224
2.0354 | -84.1208
2.0357 | -84.1194
2.036 | -84.1173
2.0363 | -84.1149
2.0366 | -84.1124
2.0369 | -84.1097
2.0372 | -84.1078
2.0376 | -84.1055
2.0379 | -84.1009
2.0382 | -84.0978
2.0385 | -84.0959
2.0388 | -84.0913
2.0391 | -84.0884
2.0394 | -84.088
2.0397 | -84.0864
2.04 | -84.0839
2.0403 | -84.0801
2.0406 | -84.0772
2.041 | -84.0755
2.0413 | -84.0724
2.0416 | -84.0645
2.0419 | -84.0597
2.0422 | -84.0584
2.0425 | -84.0575
2.0428 | -84.0554
2.0431 | -84.051
2.0434 | -84.0488
2.0437 | -84.0466
2.044 | -84.0436
2.0444 | -84.0409
2.0447 | -84.0363
2.045 | -84.0331
2.0453 | -84.0301
2.0456 | -84.0276
2.0459 | -84.0247
2.0462 | -84.025
2.0465 | -84.0245
2.0468 | -84.0197
2.0471 | -84.0173
2.0474 | -84.0147
2.0477 | -84.012
2.0481 | -84.0122
2.0484 | -84.0124
2.0487 | -84.0107
2.049 | -84.0102
2.0493 | -84.0063
2.0496 | -84.0032
2.0499 | -84.0021
2.0502 | -84.0002
2.0505 | -83.9957
2.0508 | -83.9908
2.0511 | -83.9865
2.0515 | -83.9798
2.0518 | -83.977
2.0521 | -83.9757
2.0524 | -83.9752
2.0527 | -83.9776
2.053 | -83.9768
2.0533 | -83.9749
2.0536 | -83.9723
2.0539 | -83.9703
2.0542 | -83.9675
2.0545 | -83.9643
2.0549 | -83.9626
2.0552 | -83.962
2.0555 | -83.962
2.0558 | -83.9598
2.0561 | -83.9568
2.0564 | -83.9523
2.0567 | -83.9467
2.057 | -83.9428
2.0573 | -83.9404
2.0576 | -83.9397
2.0579 | -83.9368
2.0582 | -83.9335
2.0586 | -83.9302
2.0589 | -83.927
2.0592 | -83.9246
2.0595 | -83.9221
2.0598 | -83.9204
2.0601 | -83.9189
2.0604 | -83.916
2.0607 | -83.9124
2.061 | -83.9101
2.0613 | -83.9091
2.0616 | -83.9097
2.062 | -83.9097
2.0623 | -83.9074
2.0626 | -83.9061
2.0629 | -83.9042
2.0632 | -83.9006
2.0635 | -83.8984
2.0638 | -83.8963
2.0641 | -83.8942
2.0644 | -83.892
2.0647 | -83.8887
2.065 | -83.887
2.0653 | -83.8845
2.0657 | -83.8797
2.066 | -83.8777
2.0663 | -83.876
2.0666 | -83.8764
2.0669 | -83.8753
2.0672 | -83.8745
2.0675 | -83.8722
2.0678 | -83.8711
2.0681 | -83.8712
2.0684 | -83.8702
2.0687 | -83.868
2.0691 | -83.8679
2.0694 | -83.8682
2.0697 | -83.8647
2.07 | -83.8623
2.0703 | -83.8625
2.0706 | -83.8597
2.0709 | -83.8566
2.0712 | -83.8542
2.0715 | -83.8523
2.0718 | -83.8517
2.0721 | -83.8502
2.0724 | -83.8494
2.0728 | -83.8491
2.0731 | -83.8485
2.0734 | -83.8461
2.0737 | -83.8442
2.074 | -83.8435
2.0743 | -83.8427
2.0746 | -83.8416
2.0749 | -83.8384
2.0752 | -83.8378
2.0755 | -83.839
2.0758 | -83.838
2.0762 | -83.8379
2.0765 | -83.8376
2.0768 | -83.8376
2.0771 | -83.8366
2.0774 | -83.8342
2.0777 | -83.8339
2.078 | -83.834
2.0783 | -83.834
2.0786 | -83.8332
2.0789 | -83.8325
2.0792 | -83.8321
2.0796 | -83.8331
2.0799 | -83.8338
2.0802 | -83.834
2.0805 | -83.8365
2.0808 | -83.8394
2.0811 | -83.84
2.0814 | -83.8402
2.0817 | -83.8403
2.082 | -83.84
2.0823 | -83.8407
2.0826 | -83.8414
2.0829 | -83.8411
2.0833 | -83.8408
2.0836 | -83.8409
2.0839 | -83.8415
2.0842 | -83.8415
2.0845 | -83.8434
2.0848 | -83.8432
2.0851 | -83.843
2.0854 | -83.8424
2.0857 | -83.8438
2.086 | -83.8475
2.0863 | -83.8506
2.0867 | -83.8525
2.087 | -83.8532
2.0873 | -83.854
2.0876 | -83.8541
2.0879 | -83.8555
2.0882 | -83.8576
2.0885 | -83.8585
2.0888 | -83.8544
2.0891 | -83.8501
2.0894 | -83.8489
2.0897 | -83.8507
2.09 | -83.8513
2.0904 | -83.8516
2.0907 | -83.852
2.091 | -83.8519
2.0913 | -83.8512
2.0916 | -83.8534
2.0919 | -83.8559
2.0922 | -83.8558
2.0925 | -83.8558
2.0928 | -83.8569
2.0931 | -83.8578
2.0934 | -83.8613
2.0938 | -83.8634
2.0941 | -83.8659
2.0944 | -83.868
2.0947 | -83.8713
2.095 | -83.8749
2.0953 | -83.8764
2.0956 | -83.8778
2.0959 | -83.8794
2.0962 | -83.8809
2.0965 | -83.8826
2.0968 | -83.8829
2.0972 | -83.8828
2.0975 | -83.884
2.0978 | -83.8851
2.0981 | -83.8842
2.0984 | -83.8829
2.0987 | -83.8838
2.099 | -83.8846
2.0993 | -83.886
2.0996 | -83.8871
2.0999 | -83.8882
2.1002 | -83.8886
2.1005 | -83.8894
2.1009 | -83.8907
2.1012 | -83.8941
2.1015 | -83.8965
2.1018 | -83.8977
2.1021 | -83.8992
2.1024 | -83.9008
2.1027 | -83.9042
2.103 | -83.909
2.1033 | -83.9119
2.1036 | -83.9134
2.1039 | -83.9149
2.1043 | -83.9157
2.1046 | -83.9169
2.1049 | -83.919
2.1052 | -83.9199
2.1055 | -83.9214
2.1058 | -83.9226
2.1061 | -83.9256
2.1064 | -83.9278
2.1067 | -83.9289
2.107 | -83.9284
2.1073 | -83.9303
2.1076 | -83.9325
2.108 | -83.9328
2.1083 | -83.9345
2.1086 | -83.9357
2.1089 | -83.9359
2.1092 | -83.9366
2.1095 | -83.9395
2.1098 | -83.9411
2.1101 | -83.9442
2.1104 | -83.9492
2.1107 | -83.9535
2.111 | -83.9547
2.1114 | -83.9566
2.1117 | -83.9568
2.112 | -83.9594
2.1123 | -83.9618
2.1126 | -83.9633
2.1129 | -83.965
2.1132 | -83.9669
2.1135 | -83.9689
2.1138 | -83.9714
2.1141 | -83.9723
2.1144 | -83.9751
2.1147 | -83.9801
2.1151 | -83.9826
2.1154 | -83.9846
2.1157 | -83.9858
2.116 | -83.986
2.1163 | -83.9873
2.1166 | -83.9883
2.1169 | -83.9888
2.1172 | -83.9912
2.1175 | -83.9938
2.1178 | -83.9964
2.1181 | -83.9977
2.1185 | -83.9991
2.1188 | -84.0008
2.1191 | -84.0014
2.1194 | -84.0014
2.1197 | -84.0002
2.12 | -84.001
2.1203 | -84.0029
2.1206 | -84.0049
2.1209 | -84.0061
2.1212 | -84.0067
2.1215 | -84.0059
2.1219 | -84.0042
2.1222 | -84.0051
2.1225 | -84.0066
2.1228 | -84.006
2.1231 | -84.0015
2.1234 | -84
2.1237 | -84.0007
2.124 | -84.0015
2.1243 | -84.0009
2.1246 | -83.9986
2.1249 | -84.0001
2.1252 | -84.001
2.1256 | -84.0011
2.1259 | -84.0011
2.1262 | -84.0055
2.1265 | -84.0112
2.1268 | -84.0104
2.1271 | -84.0113
2.1274 | -84.0111
2.1277 | -84.0104
2.128 | -84.0109
2.1283 | -84.0113
2.1286 | -84.0115
2.129 | -84.0119
2.1293 | -84.0127
2.1296 | -84.0136
2.1299 | -84.015
2.1302 | -84.0149
2.1305 | -84.0154
2.1308 | -84.0164
2.1311 | -84.018
2.1314 | -84.0186
2.1317 | -84.0187
2.132 | -84.018
2.1324 | -84.0196
2.1327 | -84.0213
2.133 | -84.0215
2.1333 | -84.0216
2.1336 | -84.0208
2.1339 | -84.0196
2.1342 | -84.0202
2.1345 | -84.0203
2.1348 | -84.0224
2.1351 | -84.0192
2.1354 | -84.0181
2.1357 | -84.0188
2.1361 | -84.0191
2.1364 | -84.0197
2.1367 | -84.0204
2.137 | -84.0202
2.1373 | -84.022
2.1376 | -84.0236
2.1379 | -84.0246
2.1382 | -84.026
2.1385 | -84.0287
2.1388 | -84.028
2.1391 | -84.0244
2.1395 | -84.0252
2.1398 | -84.0257
2.1401 | -84.0241
2.1404 | -84.0222
2.1407 | -84.0206
2.141 | -84.0242
2.1413 | -84.0249
2.1416 | -84.0242
2.1419 | -84.0241
2.1422 | -84.0237
2.1425 | -84.024
2.1428 | -84.0234
2.1432 | -84.0235
2.1435 | -84.0234
2.1438 | -84.0214
2.1441 | -84.0178
2.1444 | -84.0171
2.1447 | -84.0175
2.145 | -84.0172
2.1453 | -84.0182
2.1456 | -84.0185
2.1459 | -84.0168
2.1462 | -84.0147
2.1466 | -84.017
2.1469 | -84.0159
2.1472 | -84.0176
2.1475 | -84.0181
2.1478 | -84.0173
2.1481 | -84.0165
2.1484 | -84.0159
2.1487 | -84.0167
2.149 | -84.0157
2.1493 | -84.0131
2.1496 | -84.012
2.15 | -84.0116
2.1503 | -84.0095
2.1506 | -84.0063
2.1509 | -84.0051
2.1512 | -84.003
2.1515 | -84.0009
2.1518 | -83.9996
2.1521 | -83.9992
2.1524 | -83.9992
2.1527 | -83.9988
2.153 | -83.9984
2.1533 | -83.9984
2.1537 | -83.9988
2.154 | -83.9982
2.1543 | -83.9974
2.1546 | -83.9981
2.1549 | -83.9972
2.1552 | -83.9959
2.1555 | -83.9954
2.1558 | -83.9933
2.1561 | -83.993
2.1564 | -83.9937
2.1567 | -83.9931
2.1571 | -83.9905
2.1574 | -83.9893
2.1577 | -83.9892
2.158 | -83.9887
2.1583 | -83.9876
2.1586 | -83.9843
2.1589 | -83.9813
2.1592 | -83.988
2.1595 | -83.9987
2.1598 | -84.0043
2.1601 | -84.0066
2.1604 | -84.0076
2.1608 | -84.0086
2.1611 | -84.0092
2.1614 | -84.0077
2.1617 | -84.0059
2.162 | -84.0048
2.1623 | -84.0029
2.1626 | -84.0017
2.1629 | -84.0012
2.1632 | -84.0002
2.1635 | -83.999
2.1638 | -83.9977
2.1642 | -83.9968
2.1645 | -83.9934
2.1648 | -83.982
2.1651 | -83.9811
2.1654 | -83.9796
2.1657 | -83.9779
2.166 | -83.977
2.1663 | -83.9753
2.1666 | -83.9749
2.1669 | -83.975
2.1672 | -83.9741
2.1675 | -83.973
2.1679 | -83.9725
2.1682 | -83.9715
2.1685 | -83.9671
2.1688 | -83.9678
2.1691 | -83.9673
2.1694 | -83.9654
2.1697 | -83.9626
2.17 | -83.9574
2.1703 | -83.9563
2.1706 | -83.955
2.1709 | -83.9532
2.1713 | -83.952
2.1716 | -83.9501
2.1719 | -83.9484
2.1722 | -83.9496
2.1725 | -83.9468
2.1728 | -83.9418
2.1731 | -83.9417
2.1734 | -83.9394
2.1737 | -83.9393
2.174 | -83.9367
2.1743 | -83.9327
2.1747 | -83.9317
2.175 | -83.931
2.1753 | -83.9291
2.1756 | -83.9273
2.1759 | -83.9257
2.1762 | -83.9226
2.1765 | -83.9197
2.1768 | -83.9171
2.1771 | -83.9145
2.1774 | -83.9105
2.1777 | -83.9072
2.178 | -83.9052
2.1784 | -83.9028
2.1787 | -83.9008
2.179 | -83.8987
2.1793 | -83.8969
2.1796 | -83.8957
2.1799 | -83.8946
2.1802 | -83.893
2.1805 | -83.8915
2.1808 | -83.8905
2.1811 | -83.8897
2.1814 | -83.8888
2.1818 | -83.8874
2.1821 | -83.8838
2.1824 | -83.8807
2.1827 | -83.8785
2.183 | -83.8772
2.1833 | -83.8748
2.1836 | -83.8707
2.1839 | -83.8683
2.1842 | -83.8664
2.1845 | -83.8618
2.1848 | -83.8595
2.1852 | -83.8589
2.1855 | -83.858
2.1858 | -83.8594
2.1861 | -83.8598
2.1864 | -83.8576
2.1867 | -83.8552
2.187 | -83.8535
2.1873 | -83.8528
2.1876 | -83.8519
2.1879 | -83.8503
2.1882 | -83.8477
2.1885 | -83.8471
2.1889 | -83.8504
2.1892 | -83.85
2.1895 | -83.8467
2.1898 | -83.8449
2.1901 | -83.8419
2.1904 | -83.8397
2.1907 | -83.8378
2.191 | -83.8364
2.1913 | -83.8344
2.1916 | -83.8307
2.1919 | -83.8272
2.1923 | -83.8272
2.1926 | -83.8262
2.1929 | -83.8237
2.1932 | -83.8209
2.1935 | -83.818
2.1938 | -83.8156
2.1941 | -83.813
2.1944 | -83.8121
2.1947 | -83.8094
2.195 | -83.8054
2.1953 | -83.8013
2.1956 | -83.798
2.196 | -83.7905
2.1963 | -83.7859
2.1966 | -83.7827
2.1969 | -83.7803
2.1972 | -83.779
2.1975 | -83.7771
2.1978 | -83.7732
2.1981 | -83.7713
2.1984 | -83.7708
2.1987 | -83.7707
2.199 | -83.769
2.1994 | -83.769
2.1997 | -83.7668
2.2 | -83.7636
2.2003 | -83.7619
2.2006 | -83.7604
2.2009 | -83.7573
2.2012 | -83.7545
2.2015 | -83.7521
2.2018 | -83.748
2.2021 | -83.7439
2.2024 | -83.7417
2.2027 | -83.7394
2.2031 | -83.7356
2.2034 | -83.7342
2.2037 | -83.732
2.204 | -83.7293
2.2043 | -83.7274
2.2046 | -83.7275
2.2049 | -83.7266
2.2052 | -83.7243
2.2055 | -83.7215
2.2058 | -83.7189
2.2061 | -83.7163
2.2065 | -83.714
2.2068 | -83.7123
2.2071 | -83.7104
2.2074 | -83.7087
2.2077 | -83.7059
2.208 | -83.7032
2.2083 | -83.7002
2.2086 | -83.6968
2.2089 | -83.6936
2.2092 | -83.6848
2.2095 | -83.6841
2.2099 | -83.6822
2.2102 | -83.6799
2.2105 | -83.6777
2.2108 | -83.6757
2.2111 | -83.6737
2.2114 | -83.6711
2.2117 | -83.6686
2.212 | -83.6678
2.2123 | -83.6629
2.2126 | -83.6602
2.2129 | -83.6583
2.2132 | -83.655
2.2136 | -83.6526
2.2139 | -83.649
2.2142 | -83.6467
2.2145 | -83.6469
2.2148 | -83.6454
2.2151 | -83.6421
2.2154 | -83.6398
2.2157 | -83.6376
2.216 | -83.6342
2.2163 | -83.6319
2.2166 | -83.6303
2.217 | -83.6274
2.2173 | -83.6257
2.2176 | -83.6237
2.2179 | -83.6222
2.2182 | -83.622
2.2185 | -83.6199
2.2188 | -83.6189
2.2191 | -83.6194
2.2194 | -83.6197
2.2197 | -83.618
2.22 | -83.6176
2.2203 | -83.6172
2.2207 | -83.6155
2.221 | -83.6135
2.2213 | -83.6114
2.2216 | -83.6089
2.2219 | -83.6071
2.2222 | -83.6056
2.2225 | -83.6041
2.2228 | -83.6038
2.2231 | -83.6029
2.2234 | -83.601
2.2237 | -83.5994
2.2241 | -83.5978
2.2244 | -83.5964
2.2247 | -83.5945
2.225 | -83.5892
2.2253 | -83.5858
2.2256 | -83.585
2.2259 | -83.5827
2.2262 | -83.58
2.2265 | -83.5783
2.2268 | -83.5783
2.2271 | -83.5789
2.2275 | -83.5796
2.2278 | -83.5788
2.2281 | -83.578
2.2284 | -83.5771
2.2287 | -83.5759
2.229 | -83.5751
2.2293 | -83.5739
2.2296 | -83.5726
2.2299 | -83.5713
2.2302 | -83.5702
2.2305 | -83.5687
2.2308 | -83.5665
2.2312 | -83.5649
2.2315 | -83.5635
2.2318 | -83.5631
2.2321 | -83.5632
2.2324 | -83.5623
2.2327 | -83.5619
2.233 | -83.5624
2.2333 | -83.5624
2.2336 | -83.5625
2.2339 | -83.5622
2.2342 | -83.5623
2.2346 | -83.5625
2.2349 | -83.563
2.2352 | -83.5618
2.2355 | -83.5626
2.2358 | -83.5604
2.2361 | -83.5606
2.2364 | -83.5602
2.2367 | -83.561
2.237 | -83.5626
2.2373 | -83.5628
2.2376 | -83.5619
2.2379 | -83.5624
2.2383 | -83.5627
2.2386 | -83.5628
2.2389 | -83.5634
2.2392 | -83.5661
2.2395 | -83.5707
2.2398 | -83.5754
2.2401 | -83.5789
2.2404 | -83.5778
2.2407 | -83.5768
2.241 | -83.5756
2.2413 | -83.5784
2.2417 | -83.581
2.242 | -83.5807
2.2423 | -83.5807
2.2426 | -83.5809
2.2429 | -83.5827
2.2432 | -83.5843
2.2435 | -83.5846
2.2438 | -83.5851
2.2441 | -83.5859
2.2444 | -83.5864
2.2447 | -83.5896
2.245 | -83.5907
2.2454 | -83.5926
2.2457 | -83.594
2.246 | -83.5946
2.2463 | -83.5951
2.2466 | -83.5971
2.2469 | -83.5992
2.2472 | -83.6005
2.2475 | -83.6016
2.2478 | -83.6037
2.2481 | -83.6067
2.2484 | -83.6094
2.2488 | -83.611
2.2491 | -83.6126
2.2494 | -83.6163
2.2497 | -83.6199
2.25 | -83.622
2.2503 | -83.6238
2.2506 | -83.6257
2.2509 | -83.6246
2.2512 | -83.6236
2.2515 | -83.6247
2.2518 | -83.6258
2.2522 | -83.6277
2.2525 | -83.63
2.2528 | -83.6317
2.2531 | -83.634
2.2534 | -83.6363
2.2537 | -83.6417
2.254 | -83.6447
2.2543 | -83.6474
2.2546 | -83.6487
2.2549 | -83.6506
2.2552 | -83.653
2.2555 | -83.6555
2.2559 | -83.6582
2.2562 | -83.6615
2.2565 | -83.6643
2.2568 | -83.6672
2.2571 | -83.6716
2.2574 | -83.6752
2.2577 | -83.6785
2.258 | -83.6817
2.2583 | -83.6846
2.2586 | -83.6875
2.2589 | -83.691
2.2593 | -83.6946
2.2596 | -83.6971
2.2599 | -83.6995
2.2602 | -83.7029
2.2605 | -83.706
2.2608 | -83.7107
2.2611 | -83.7158
2.2614 | -83.7195
2.2617 | -83.7229
2.262 | -83.7259
2.2623 | -83.7285
2.2626 | -83.7324
2.263 | -83.7337
2.2633 | -83.7342
2.2636 | -83.7357
2.2639 | -83.7392
2.2642 | -83.7436
2.2645 | -83.7469
2.2648 | -83.7506
2.2651 | -83.7534
2.2654 | -83.7553
2.2657 | -83.7582
2.266 | -83.7613
2.2664 | -83.7652
2.2667 | -83.7697
2.267 | -83.7739
2.2673 | -83.7782
2.2676 | -83.782
2.2679 | -83.7858
2.2682 | -83.7915
2.2685 | -83.7976
2.2688 | -83.8019
2.2691 | -83.8063
2.2694 | -83.8095
2.2698 | -83.8137
2.2701 | -83.8176
2.2704 | -83.82
2.2707 | -83.8222
2.271 | -83.8265
2.2713 | -83.8301
2.2716 | -83.8348
2.2719 | -83.8398
2.2722 | -83.8455
2.2725 | -83.8524
2.2728 | -83.858
2.2731 | -83.8619
2.2735 | -83.866
2.2738 | -83.8701
2.2741 | -83.8749
2.2744 | -83.8799
2.2747 | -83.8871
2.275 | -83.8946
2.2753 | -83.8992
2.2756 | -83.9036
2.2759 | -83.9104
2.2762 | -83.9201
2.2765 | -83.9262
2.2769 | -83.9297
2.2772 | -83.9347
2.2775 | -83.9402
2.2778 | -83.9462
2.2781 | -83.9509
2.2784 | -83.9554
2.2787 | -83.9602
2.279 | -83.9644
2.2793 | -83.9664
2.2796 | -83.9715
2.2799 | -83.9784
2.2803 | -83.9831
2.2806 | -83.9882
2.2809 | -83.994
2.2812 | -83.9983
2.2815 | -84.0025
2.2818 | -84.0073
2.2821 | -84.0095
2.2824 | -84.0124
2.2827 | -84.0175
2.283 | -84.0202
2.2833 | -84.0215
2.2836 | -84.0255
2.284 | -84.0306
2.2843 | -84.0353
2.2846 | -84.0401
2.2849 | -84.0454
2.2852 | -84.05
2.2855 | -84.0552
2.2858 | -84.0607
2.2861 | -84.0652
2.2864 | -84.0693
2.2867 | -84.0725
2.287 | -84.0751
2.2874 | -84.0808
2.2877 | -84.0887
2.288 | -84.094
2.2883 | -84.0986
2.2886 | -84.1013
2.2889 | -84.1091
2.2892 | -84.1149
2.2895 | -84.1201
2.2898 | -84.124
2.2901 | -84.1264
2.2904 | -84.1265
2.2907 | -84.1305
2.2911 | -84.1363
2.2914 | -84.141
2.2917 | -84.1448
2.292 | -84.1492
2.2923 | -84.1537
2.2926 | -84.1585
2.2929 | -84.1629
2.2932 | -84.1685
2.2935 | -84.1764
2.2938 | -84.1815
2.2941 | -84.184
2.2945 | -84.1845
2.2948 | -84.1891
2.2951 | -84.1964
2.2954 | -84.2002
2.2957 | -84.2023
2.296 | -84.2052
2.2963 | -84.209
2.2966 | -84.2134
2.2969 | -84.2168
2.2972 | -84.2189
2.2975 | -84.2217
2.2978 | -84.2257
2.2982 | -84.2272
2.2985 | -84.2324
2.2988 | -84.2389
2.2991 | -84.2438
2.2994 | -84.245
2.2997 | -84.246
2.3 | -84.2491
2.3003 | -84.2531
2.3006 | -84.259
2.3009 | -84.2641
2.3012 | -84.2675
2.3016 | -84.2713
2.3019 | -84.2754
2.3022 | -84.2794
2.3025 | -84.2845
2.3028 | -84.2863
2.3031 | -84.2859
2.3034 | -84.2896
2.3037 | -84.293
2.304 | -84.2973
2.3043 | -84.3023
2.3046 | -84.3044
2.305 | -84.307
2.3053 | -84.311
2.3056 | -84.3155
2.3059 | -84.3199
2.3062 | -84.3245
2.3065 | -84.328
2.3068 | -84.3302
2.3071 | -84.3316
2.3074 | -84.3337
2.3077 | -84.3335
2.308 | -84.3347
2.3083 | -84.3369
2.3087 | -84.3391
2.309 | -84.3412
2.3093 | -84.343
2.3096 | -84.3446
2.3099 | -84.3472
2.3102 | -84.35
2.3105 | -84.3529
2.3108 | -84.3545
2.3111 | -84.3569
2.3114 | -84.3601
2.3117 | -84.3623
2.3121 | -84.3645
2.3124 | -84.3673
2.3127 | -84.3705
2.313 | -84.3722
2.3133 | -84.3717
2.3136 | -84.3729
2.3139 | -84.3734
2.3142 | -84.3764
2.3145 | -84.3805
2.3148 | -84.3833
2.3151 | -84.3843
2.3154 | -84.3854
2.3158 | -84.3863
2.3161 | -84.3872
2.3164 | -84.3887
2.3167 | -84.3901
2.317 | -84.3915
2.3173 | -84.3928
2.3176 | -84.3939
2.3179 | -84.3949
2.3182 | -84.3967
2.3185 | -84.3978
2.3188 | -84.3988
2.3192 | -84.4004
2.3195 | -84.4021
2.3198 | -84.4047
2.3201 | -84.4036
2.3204 | -84.4045
2.3207 | -84.4128
2.321 | -84.4166
2.3213 | -84.4169
2.3216 | -84.4178
2.3219 | -84.4177
2.3222 | -84.4184
2.3226 | -84.4198
2.3229 | -84.4204
2.3232 | -84.4214
2.3235 | -84.4219
2.3238 | -84.4223
2.3241 | -84.4208
2.3244 | -84.4209
2.3247 | -84.4244
2.325 | -84.4232
2.3253 | -84.423
2.3256 | -84.4241
2.3259 | -84.4253
2.3263 | -84.428
2.3266 | -84.4333
2.3269 | -84.4351
2.3272 | -84.4347
2.3275 | -84.4341
2.3278 | -84.4334
2.3281 | -84.4333
2.3284 | -84.4333
2.3287 | -84.4345
2.329 | -84.4381
2.3293 | -84.4419
2.3297 | -84.4418
2.33 | -84.4425
2.3303 | -84.443
2.3306 | -84.4424
2.3309 | -84.4432
2.3312 | -84.4455
2.3315 | -84.4459
2.3318 | -84.4462
2.3321 | -84.4478
2.3324 | -84.448
2.3327 | -84.4471
2.333 | -84.4481
2.3334 | -84.4446
2.3337 | -84.4417
2.334 | -84.4402
2.3343 | -84.4403
2.3346 | -84.4392
2.3349 | -84.4357
2.3352 | -84.4325
2.3355 | -84.4306
2.3358 | -84.4291
2.3361 | -84.4284
2.3364 | -84.4273
2.3368 | -84.4256
2.3371 | -84.4233
2.3374 | -84.4217
2.3377 | -84.4201
2.338 | -84.4186
2.3383 | -84.4176
2.3386 | -84.4145
2.3389 | -84.4121
2.3392 | -84.4127
2.3395 | -84.4132
2.3398 | -84.4127
2.3402 | -84.4119
2.3405 | -84.4084
2.3408 | -84.4072
2.3411 | -84.4085
2.3414 | -84.4086
2.3417 | -84.4064
2.342 | -84.4071
2.3423 | -84.4074
2.3426 | -84.4039
2.3429 | -84.4023
2.3432 | -84.3997
2.3435 | -84.3964
2.3439 | -84.3948
2.3442 | -84.3933
2.3445 | -84.3926
2.3448 | -84.3908
2.3451 | -84.389
2.3454 | -84.387
2.3457 | -84.384
2.346 | -84.3833
2.3463 | -84.3823
2.3466 | -84.381
2.3469 | -84.3792
2.3473 | -84.3782
2.3476 | -84.3765
2.3479 | -84.377
2.3482 | -84.3777
2.3485 | -84.3755
2.3488 | -84.3708
2.3491 | -84.3665
2.3494 | -84.3644
2.3497 | -84.3652
2.35 | -84.3652
2.3503 | -84.364
2.3506 | -84.3629
2.351 | -84.3622
2.3513 | -84.3607
2.3516 | -84.3594
2.3519 | -84.3595
2.3522 | -84.3572
2.3525 | -84.3563
2.3528 | -84.3549
2.3531 | -84.3531
2.3534 | -84.3547
2.3537 | -84.3546
2.354 | -84.3546
2.3544 | -84.3591
2.3547 | -84.3583
2.355 | -84.3492
2.3553 | -84.344
2.3556 | -84.342
2.3559 | -84.3402
2.3562 | -84.3401
2.3565 | -84.338
2.3568 | -84.3371
2.3571 | -84.3361
2.3574 | -84.3358
2.3578 | -84.3355
2.3581 | -84.3356
2.3584 | -84.3352
2.3587 | -84.3333
2.359 | -84.3326
2.3593 | -84.3298
2.3596 | -84.328
2.3599 | -84.3268
2.3602 | -84.3243
2.3605 | -84.321
2.3608 | -84.3218
2.3611 | -84.3222
2.3615 | -84.3205
2.3618 | -84.3195
2.3621 | -84.3192
2.3624 | -84.3172
2.3627 | -84.3154
2.363 | -84.3115
2.3633 | -84.3083
2.3636 | -84.3065
2.3639 | -84.3062
2.3642 | -84.3062
2.3645 | -84.3043
2.3649 | -84.3037
2.3652 | -84.3029
2.3655 | -84.3019
2.3658 | -84.3002
2.3661 | -84.299
2.3664 | -84.2968
2.3667 | -84.2961
2.367 | -84.296
2.3673 | -84.2958
2.3676 | -84.2959
2.3679 | -84.2966
2.3682 | -84.3022
2.3686 | -84.3034
2.3689 | -84.3025
2.3692 | -84.3019
2.3695 | -84.3035
2.3698 | -84.3031
2.3701 | -84.3025
2.3704 | -84.3009
2.3707 | -84.299
2.371 | -84.2984
2.3713 | -84.2949
2.3716 | -84.2916
2.372 | -84.2934
2.3723 | -84.2919
2.3726 | -84.2908
2.3729 | -84.2915
2.3732 | -84.2918
2.3735 | -84.2918
2.3738 | -84.2911
2.3741 | -84.2904
2.3744 | -84.2897
2.3747 | -84.2889
2.375 | -84.287
2.3754 | -84.286
2.3757 | -84.2845
2.376 | -84.2838
2.3763 | -84.283
2.3766 | -84.2816
2.3769 | -84.2781
2.3772 | -84.2772
2.3775 | -84.2764
2.3778 | -84.2757
2.3781 | -84.2752
2.3784 | -84.2731
2.3787 | -84.2729
2.3791 | -84.2717
2.3794 | -84.2693
2.3797 | -84.2672
2.38 | -84.267
2.3803 | -84.2662
2.3806 | -84.265
2.3809 | -84.2661
2.3812 | -84.2653
2.3815 | -84.2637
2.3818 | -84.2639
2.3821 | -84.2636
2.3825 | -84.2626
2.3828 | -84.2624
2.3831 | -84.2624
2.3834 | -84.2627
2.3837 | -84.2619
2.384 | -84.2599
2.3843 | -84.2607
2.3846 | -84.2627
2.3849 | -84.2632
2.3852 | -84.2664
2.3855 | -84.2682
2.3858 | -84.2683
2.3862 | -84.2661
2.3865 | -84.2606
2.3868 | -84.2567
2.3871 | -84.2549
2.3874 | -84.2546
2.3877 | -84.2532
2.388 | -84.2527
2.3883 | -84.2521
2.3886 | -84.2523
2.3889 | -84.2511
2.3892 | -84.2488
2.3896 | -84.2477
2.3899 | -84.2461
2.3902 | -84.2443
2.3905 | -84.2443
2.3908 | -84.2438
2.3911 | -84.2418
2.3914 | -84.2404
2.3917 | -84.2405
2.392 | -84.2414
2.3923 | -84.2414
2.3926 | -84.24
2.3929 | -84.2379
2.3933 | -84.2378
2.3936 | -84.2371
2.3939 | -84.2351
2.3942 | -84.2344
2.3945 | -84.2345
2.3948 | -84.234
2.3951 | -84.234
2.3954 | -84.2331
2.3957 | -84.2322
2.396 | -84.234
2.3963 | -84.2347
2.3967 | -84.2329
2.397 | -84.2324
2.3973 | -84.2317
2.3976 | -84.2288
2.3979 | -84.2279
2.3982 | -84.2269
2.3985 | -84.2257
2.3988 | -84.225
2.3991 | -84.2254
2.3994 | -84.224
2.3997 | -84.2232
2.4001 | -84.2225
2.4004 | -84.2216
2.4007 | -84.2177
2.401 | -84.2153
2.4013 | -84.2144
2.4016 | -84.2138
2.4019 | -84.2154
2.4022 | -84.2153
2.4025 | -84.2141
2.4028 | -84.2129
2.4031 | -84.2111
2.4034 | -84.2113
2.4038 | -84.2121
2.4041 | -84.2109
2.4044 | -84.2088
2.4047 | -84.206
2.405 | -84.2044
2.4053 | -84.2013
2.4056 | -84.1996
2.4059 | -84.1976
2.4062 | -84.1955
2.4065 | -84.1941
2.4068 | -84.1919
2.4072 | -84.1886
2.4075 | -84.1863
2.4078 | -84.1851
2.4081 | -84.1826
2.4084 | -84.1772
2.4087 | -84.1756
2.409 | -84.171
2.4093 | -84.1699
2.4096 | -84.1668
2.4099 | -84.1657
2.4102 | -84.1645
2.4106 | -84.1624
2.4109 | -84.1589
2.4112 | -84.1541
2.4115 | -84.1515
2.4118 | -84.1501
2.4121 | -84.1488
2.4124 | -84.1459
2.4127 | -84.142
2.413 | -84.1418
2.4133 | -84.1411
2.4136 | -84.1402
2.4139 | -84.1346
2.4143 | -84.1224
2.4146 | -84.12
2.4149 | -84.117
2.4152 | -84.1155
2.4155 | -84.1133
2.4158 | -84.1113
2.4161 | -84.1074
2.4164 | -84.1045
2.4167 | -84.102
2.417 | -84.0995
2.4173 | -84.0988
2.4177 | -84.0968
2.418 | -84.0943
2.4183 | -84.0918
2.4186 | -84.0882
2.4189 | -84.0818
2.4192 | -84.0791
2.4195 | -84.0765
2.4198 | -84.0731
2.4201 | -84.0702
2.4204 | -84.0671
2.4207 | -84.0632
2.421 | -84.0604
2.4214 | -84.0574
2.4217 | -84.0565
2.422 | -84.0535
2.4223 | -84.0516
2.4226 | -84.0503
2.4229 | -84.0474
2.4232 | -84.0444
2.4235 | -84.042
2.4238 | -84.039
2.4241 | -84.0363
2.4244 | -84.0338
2.4248 | -84.0308
2.4251 | -84.0283
2.4254 | -84.0262
2.4257 | -84.025
2.426 | -84.0205
2.4263 | -84.0157
2.4266 | -84.0137
2.4269 | -84.0126
2.4272 | -84.0095
2.4275 | -84.0054
2.4278 | -84.0022
2.4282 | -83.9985
2.4285 | -83.995
2.4288 | -83.9929
2.4291 | -83.9904
2.4294 | -83.9863
2.4297 | -83.9821
2.43 | -83.9785
2.4303 | -83.9747
2.4306 | -83.9707
2.4309 | -83.9683
2.4312 | -83.9652
2.4315 | -83.9607
2.4319 | -83.9542
2.4322 | -83.9506
2.4325 | -83.9476
2.4328 | -83.9446
2.4331 | -83.942
2.4334 | -83.9363
2.4337 | -83.933
2.434 | -83.9304
2.4343 | -83.9285
2.4346 | -83.9226
2.4349 | -83.9166
2.4353 | -83.914
2.4356 | -83.9102
2.4359 | -83.9058
2.4362 | -83.9018
2.4365 | -83.8964
2.4368 | -83.8916
2.4371 | -83.8884
2.4374 | -83.8841
2.4377 | -83.8807
2.438 | -83.8766
2.4383 | -83.8731
2.4386 | -83.8683
2.439 | -83.8597
2.4393 | -83.8537
2.4396 | -83.8482
2.4399 | -83.8439
2.4402 | -83.8396
2.4405 | -83.8353
2.4408 | -83.8294
2.4411 | -83.8237
2.4414 | -83.8181
2.4417 | -83.8135
2.442 | -83.8095
2.4424 | -83.8042
2.4427 | -83.8
2.443 | -83.7965
2.4433 | -83.7933
2.4436 | -83.7913
2.4439 | -83.7872
2.4442 | -83.7824
2.4445 | -83.7768
2.4448 | -83.7701
2.4451 | -83.7655
2.4454 | -83.761
2.4457 | -83.7568
2.4461 | -83.7526
2.4464 | -83.7471
2.4467 | -83.7437
2.447 | -83.7413
2.4473 | -83.7355
2.4476 | -83.7329
2.4479 | -83.7301
2.4482 | -83.7209
2.4485 | -83.7117
2.4488 | -83.7048
2.4491 | -83.7002
2.4495 | -83.6947
2.4498 | -83.6905
2.4501 | -83.6859
2.4504 | -83.6806
2.4507 | -83.6761
2.451 | -83.6715
2.4513 | -83.6656
2.4516 | -83.6612
2.4519 | -83.6573
2.4522 | -83.6513
2.4525 | -83.6457
2.4529 | -83.641
2.4532 | -83.637
2.4535 | -83.6341
2.4538 | -83.6299
2.4541 | -83.6252
2.4544 | -83.6204
2.4547 | -83.6148
2.455 | -83.6102
2.4553 | -83.6069
2.4556 | -83.6046
2.4559 | -83.5994
2.4562 | -83.5941
2.4566 | -83.5893
2.4569 | -83.5839
2.4572 | -83.5796
2.4575 | -83.5752
2.4578 | -83.5706
2.4581 | -83.5668
2.4584 | -83.5623
2.4587 | -83.5591
2.459 | -83.5546
2.4593 | -83.5499
2.4596 | -83.5475
2.46 | -83.5476
2.4603 | -83.5446
2.4606 | -83.5402
2.4609 | -83.5363
2.4612 | -83.5334
2.4615 | -83.5298
2.4618 | -83.5251
2.4621 | -83.52
2.4624 | -83.5162
2.4627 | -83.5116
2.463 | -83.5059
2.4633 | -83.5016
2.4637 | -83.4979
2.464 | -83.4944
2.4643 | -83.4907
2.4646 | -83.486
2.4649 | -83.4821
2.4652 | -83.4798
2.4655 | -83.4769
2.4658 | -83.4731
2.4661 | -83.4695
2.4664 | -83.4646
2.4667 | -83.4582
2.4671 | -83.4538
2.4674 | -83.4494
2.4677 | -83.4456
2.468 | -83.442
2.4683 | -83.4385
2.4686 | -83.4344
2.4689 | -83.4309
2.4692 | -83.4295
2.4695 | -83.4258
2.4698 | -83.422
2.4701 | -83.4181
2.4705 | -83.4139
2.4708 | -83.4094
2.4711 | -83.4044
2.4714 | -83.4009
2.4717 | -83.3977
2.472 | -83.3941
2.4723 | -83.39
2.4726 | -83.3865
2.4729 | -83.3837
2.4732 | -83.3796
2.4735 | -83.3746
2.4738 | -83.3716
2.4742 | -83.3703
2.4745 | -83.3678
2.4748 | -83.3645
2.4751 | -83.3616
2.4754 | -83.3586
2.4757 | -83.3547
2.476 | -83.3516
2.4763 | -83.3508
2.4766 | -83.3462
2.4769 | -83.3428
2.4772 | -83.3398
2.4776 | -83.3368
2.4779 | -83.3326
2.4782 | -83.3292
2.4785 | -83.3258
2.4788 | -83.3211
2.4791 | -83.3186
2.4794 | -83.3161
2.4797 | -83.3137
2.48 | -83.3132
2.4803 | -83.3116
2.4806 | -83.3102
2.4809 | -83.3092
2.4813 | -83.3062
2.4816 | -83.3031
2.4819 | -83.3011
2.4822 | -83.2988
2.4825 | -83.2963
2.4828 | -83.294
2.4831 | -83.292
2.4834 | -83.29
2.4837 | -83.2878
2.484 | -83.2861
2.4843 | -83.284
2.4847 | -83.2823
2.485 | -83.2802
2.4853 | -83.2817
2.4856 | -83.2803
2.4859 | -83.2792
2.4862 | -83.2768
2.4865 | -83.2744
2.4868 | -83.2723
2.4871 | -83.2701
2.4874 | -83.2696
2.4877 | -83.269
2.4881 | -83.2672
2.4884 | -83.2662
2.4887 | -83.2642
2.489 | -83.2611
2.4893 | -83.2589
2.4896 | -83.257
2.4899 | -83.2553
2.4902 | -83.2542
2.4905 | -83.2533
2.4908 | -83.2524
2.4911 | -83.2516
2.4914 | -83.2517
2.4918 | -83.2522
2.4921 | -83.2517
2.4924 | -83.2509
2.4927 | -83.2507
2.493 | -83.2499
2.4933 | -83.2487
2.4936 | -83.2471
2.4939 | -83.2456
2.4942 | -83.2441
2.4945 | -83.2439
2.4948 | -83.2444
2.4952 | -83.2446
2.4955 | -83.2433
2.4958 | -83.2428
2.4961 | -83.243
2.4964 | -83.2431
2.4967 | -83.2422
2.497 | -83.2407
2.4973 | -83.2382
2.4976 | -83.2358
2.4979 | -83.2346
2.4982 | -83.2343
2.4985 | -83.2349
2.4989 | -83.234
2.4992 | -83.2334
2.4995 | -83.2332
2.4998 | -83.2304
2.5001 | -83.2301
2.5004 | -83.2304
2.5007 | -83.2303
2.501 | -83.2308
2.5013 | -83.2341
2.5016 | -83.2368
2.5019 | -83.2369
2.5023 | -83.2354
2.5026 | -83.2346
2.5029 | -83.2367
2.5032 | -83.24
2.5035 | -83.2424
2.5038 | -83.2422
2.5041 | -83.2434
2.5044 | -83.2443
2.5047 | -83.2424
2.505 | -83.2411
2.5053 | -83.2415
2.5057 | -83.2401
2.506 | -83.2403
2.5063 | -83.2411
2.5066 | -83.2416
2.5069 | -83.2409
2.5072 | -83.2431
2.5075 | -83.2452
2.5078 | -83.2455
2.5081 | -83.2459
2.5084 | -83.2463
2.5087 | -83.2458
2.509 | -83.2478
2.5094 | -83.2479
2.5097 | -83.248
2.51 | -83.2475
2.5103 | -83.2491
2.5106 | -83.2511
2.5109 | -83.2522
2.5112 | -83.2546
2.5115 | -83.2562
2.5118 | -83.256
2.5121 | -83.2564
2.5124 | -83.2579
2.5128 | -83.258
2.5131 | -83.2557
2.5134 | -83.2564
2.5137 | -83.2576
2.514 | -83.2589
2.5143 | -83.2604
2.5146 | -83.262
2.5149 | -83.2619
2.5152 | -83.2612
2.5155 | -83.2617
2.5158 | -83.263
2.5161 | -83.2641
2.5165 | -83.2655
2.5168 | -83.2658
2.5171 | -83.2674
2.5174 | -83.267
2.5177 | -83.2687
2.518 | -83.2725
2.5183 | -83.276
2.5186 | -83.2783
2.5189 | -83.2791
2.5192 | -83.2808
2.5195 | -83.279
2.5199 | -83.2782
2.5202 | -83.28
2.5205 | -83.2823
2.5208 | -83.2843
2.5211 | -83.2872
2.5214 | -83.2892
2.5217 | -83.291
2.522 | -83.2937
2.5223 | -83.2963
2.5226 | -83.2985
2.5229 | -83.3011
2.5232 | -83.3027
2.5236 | -83.305
2.5239 | -83.3041
2.5242 | -83.3049
2.5245 | -83.3081
2.5248 | -83.3155
2.5251 | -83.3182
2.5254 | -83.3177
2.5257 | -83.3194
2.526 | -83.32
2.5263 | -83.3211
2.5266 | -83.3224
2.527 | -83.3212
2.5273 | -83.3216
2.5276 | -83.3224
2.5279 | -83.3237
2.5282 | -83.3249
2.5285 | -83.3269
2.5288 | -83.328
2.5291 | -83.3308
2.5294 | -83.3337
2.5297 | -83.3345
2.53 | -83.334
2.5304 | -83.3356
2.5307 | -83.3387
2.531 | -83.3394
2.5313 | -83.3394
2.5316 | -83.341
2.5319 | -83.3426
2.5322 | -83.3443
2.5325 | -83.3455
2.5328 | -83.3476
2.5331 | -83.3499
2.5334 | -83.353
2.5337 | -83.3533
2.5341 | -83.3524
2.5344 | -83.3522
2.5347 | -83.3531
2.535 | -83.355
2.5353 | -83.3531
2.5356 | -83.3526
2.5359 | -83.3536
2.5362 | -83.3552
2.5365 | -83.3538
2.5368 | -83.3555
2.5371 | -83.3573
2.5375 | -83.3587
2.5378 | -83.3588
2.5381 | -83.3583
2.5384 | -83.359
2.5387 | -83.3601
2.539 | -83.361
2.5393 | -83.3615
2.5396 | -83.3629
2.5399 | -83.3641
2.5402 | -83.365
2.5405 | -83.3662
2.5408 | -83.367
2.5412 | -83.3684
2.5415 | -83.3679
2.5418 | -83.3683
2.5421 | -83.3699
2.5424 | -83.3737
2.5427 | -83.3763
2.543 | -83.3775
2.5433 | -83.3799
2.5436 | -83.3814
2.5439 | -83.3821
2.5442 | -83.3829
2.5446 | -83.3835
2.5449 | -83.3841
2.5452 | -83.3862
2.5455 | -83.3869
2.5458 | -83.3865
2.5461 | -83.387
2.5464 | -83.3884
2.5467 | -83.3901
2.547 | -83.3908
2.5473 | -83.3918
2.5476 | -83.3919
2.548 | -83.393
2.5483 | -83.3956
2.5486 | -83.3973
2.5489 | -83.3982
2.5492 | -83.3978
2.5495 | -83.3969
2.5498 | -83.3987
2.5501 | -83.4006
2.5504 | -83.4001
2.5507 | -83.4001
2.551 | -83.4016
2.5513 | -83.4033
2.5517 | -83.4033
2.552 | -83.4038
2.5523 | -83.405
2.5526 | -83.4055
2.5529 | -83.4052
2.5532 | -83.4024
2.5535 | -83.4015
2.5538 | -83.401
2.5541 | -83.4016
2.5544 | -83.3991
2.5547 | -83.3975
2.5551 | -83.3974
2.5554 | -83.3982
2.5557 | -83.3995
2.556 | -83.3997
2.5563 | -83.4
2.5566 | -83.3967
2.5569 | -83.3945
2.5572 | -83.3954
2.5575 | -83.3957
2.5578 | -83.3954
2.5581 | -83.3974
2.5585 | -83.4019
2.5588 | -83.4028
2.5591 | -83.4012
2.5594 | -83.3999
2.5597 | -83.3999
2.56 | -83.4003
2.5603 | -83.3975
2.5606 | -83.3929
2.5609 | -83.3912
2.5612 | -83.39
2.5615 | -83.3888
2.5618 | -83.3891
2.5622 | -83.3892
2.5625 | -83.3893
2.5628 | -83.3887
2.5631 | -83.3874
2.5634 | -83.3852
2.5637 | -83.3806
2.564 | -83.3786
2.5643 | -83.377
2.5646 | -83.374
2.5649 | -83.3729
2.5652 | -83.3729
2.5656 | -83.3711
2.5659 | -83.3684
2.5662 | -83.3662
2.5665 | -83.3659
2.5668 | -83.3616
2.5671 | -83.3546
2.5674 | -83.352
2.5677 | -83.3501
2.568 | -83.3484
2.5683 | -83.3475
2.5686 | -83.3446
2.5689 | -83.3426
2.5693 | -83.3402
2.5696 | -83.3388
2.5699 | -83.338
2.5702 | -83.3387
2.5705 | -83.3391
2.5708 | -83.3364
2.5711 | -83.3346
2.5714 | -83.3308
2.5717 | -83.3268
2.572 | -83.3214
2.5723 | -83.3186
2.5727 | -83.3168
2.573 | -83.3138
2.5733 | -83.309
2.5736 | -83.3074
2.5739 | -83.3059
2.5742 | -83.2995
2.5745 | -83.2939
2.5748 | -83.2909
2.5751 | -83.2882
2.5754 | -83.2864
2.5757 | -83.2867
2.576 | -83.2844
2.5764 | -83.2817
2.5767 | -83.277
2.577 | -83.2716
2.5773 | -83.269
2.5776 | -83.2667
2.5779 | -83.2653
2.5782 | -83.2632
2.5785 | -83.2606
2.5788 | -83.2579
2.5791 | -83.2554
2.5794 | -83.2532
2.5798 | -83.2513
2.5801 | -83.2493
2.5804 | -83.2474
2.5807 | -83.2461
2.581 | -83.2443
2.5813 | -83.2406
2.5816 | -83.2372
2.5819 | -83.236
2.5822 | -83.233
2.5825 | -83.2321
2.5828 | -83.2278
2.5832 | -83.2239
2.5835 | -83.2164
2.5838 | -83.2134
2.5841 | -83.2092
2.5844 | -83.2057
2.5847 | -83.2026
2.585 | -83.199
2.5853 | -83.1957
2.5856 | -83.1925
2.5859 | -83.1888
2.5862 | -83.1868
2.5865 | -83.1843
2.5869 | -83.1812
2.5872 | -83.1781
2.5875 | -83.1723
2.5878 | -83.1711
2.5881 | -83.1722
2.5884 | -83.1709
2.5887 | -83.1657
2.589 | -83.1604
2.5893 | -83.1544
2.5896 | -83.1489
2.5899 | -83.1441
2.5903 | -83.1415
2.5906 | -83.1379
2.5909 | -83.1339
2.5912 | -83.1302
2.5915 | -83.1248
2.5918 | -83.12
2.5921 | -83.1163
2.5924 | -83.113
2.5927 | -83.109
2.593 | -83.1049
2.5933 | -83.1027
2.5936 | -83.0984
2.594 | -83.0939
2.5943 | -83.0908
2.5946 | -83.0848
2.5949 | -83.0792
2.5952 | -83.0763
2.5955 | -83.0726
2.5958 | -83.0689
2.5961 | -83.0621
2.5964 | -83.0557
2.5967 | -83.0494
2.597 | -83.0448
2.5974 | -83.0408
2.5977 | -83.0371
2.598 | -83.0323
2.5983 | -83.0281
2.5986 | -83.0234
2.5989 | -83.0183
2.5992 | -83.0145
2.5995 | -83.0105
2.5998 | -83.0068
2.6001 | -83.0033
2.6004 | -82.9978
2.6008 | -82.993
2.6011 | -82.9888
2.6014 | -82.9868
2.6017 | -82.9816
2.602 | -82.9739
2.6023 | -82.9686
2.6026 | -82.9652
2.6029 | -82.9608
2.6032 | -82.9569
2.6035 | -82.9512
2.6038 | -82.946
2.6041 | -82.9418
2.6045 | -82.9379
2.6048 | -82.9328
2.6051 | -82.9275
2.6054 | -82.9233
2.6057 | -82.9182
2.606 | -82.9111
2.6063 | -82.9065
2.6066 | -82.9024
2.6069 | -82.8994
2.6072 | -82.8953
2.6075 | -82.8905
2.6079 | -82.8854
2.6082 | -82.8806
2.6085 | -82.8757
2.6088 | -82.8728
2.6091 | -82.8686
2.6094 | -82.8649
2.6097 | -82.8585
2.61 | -82.8537
2.6103 | -82.85
2.6106 | -82.8473
2.6109 | -82.8424
2.6112 | -82.8367
2.6116 | -82.8323
2.6119 | -82.8277
2.6122 | -82.8244
2.6125 | -82.8205
2.6128 | -82.816
2.6131 | -82.8125
2.6134 | -82.8081
2.6137 | -82.804
2.614 | -82.7998
2.6143 | -82.7966
2.6146 | -82.7953
2.615 | -82.7933
2.6153 | -82.7888
2.6156 | -82.7832
2.6159 | -82.7779
2.6162 | -82.7729
2.6165 | -82.7686
2.6168 | -82.7649
2.6171 | -82.7605
2.6174 | -82.755
2.6177 | -82.7489
2.618 | -82.7455
2.6184 | -82.7405
2.6187 | -82.7344
2.619 | -82.7307
2.6193 | -82.7265
2.6196 | -82.7223
2.6199 | -82.718
2.6202 | -82.7133
2.6205 | -82.709
2.6208 | -82.7062
2.6211 | -82.7041
2.6214 | -82.7006
2.6217 | -82.6966
2.6221 | -82.6935
2.6224 | -82.6892
2.6227 | -82.6854
2.623 | -82.6821
2.6233 | -82.6778
2.6236 | -82.6746
2.6239 | -82.6679
2.6242 | -82.662
2.6245 | -82.6598
2.6248 | -82.6577
2.6251 | -82.6526
2.6255 | -82.6462
2.6258 | -82.6415
2.6261 | -82.6375
2.6264 | -82.6329
2.6267 | -82.6251
2.627 | -82.6171
2.6273 | -82.6099
2.6276 | -82.6042
2.6279 | -82.6005
2.6282 | -82.5974
2.6285 | -82.5938
2.6288 | -82.5895
2.6292 | -82.5862
2.6295 | -82.5818
2.6298 | -82.5753
2.6301 | -82.5696
2.6304 | -82.5633
2.6307 | -82.5585
2.631 | -82.5543
2.6313 | -82.55
2.6316 | -82.5455
2.6319 | -82.5416
2.6322 | -82.5371
2.6326 | -82.5302
2.6329 | -82.5238
2.6332 | -82.5179
2.6335 | -82.5115
2.6338 | -82.5083
2.6341 | -82.5017
2.6344 | -82.4973
2.6347 | -82.492
2.635 | -82.4869
2.6353 | -82.4822
2.6356 | -82.476
2.6359 | -82.4721
2.6363 | -82.4669
2.6366 | -82.4618
2.6369 | -82.4589
2.6372 | -82.4547
2.6375 | -82.4505
2.6378 | -82.4459
2.6381 | -82.4416
2.6384 | -82.4377
2.6387 | -82.4336
2.639 | -82.4298
2.6393 | -82.4243
2.6397 | -82.4201
2.64 | -82.416
2.6403 | -82.4129
2.6406 | -82.41
2.6409 | -82.4042
2.6412 | -82.3968
2.6415 | -82.3903
2.6418 | -82.3854
2.6421 | -82.3821
2.6424 | -82.3757
2.6427 | -82.3699
2.6431 | -82.3661
2.6434 | -82.3617
2.6437 | -82.3567
2.644 | -82.3521
2.6443 | -82.3502
2.6446 | -82.3457
2.6449 | -82.3417
2.6452 | -82.3382
2.6455 | -82.334
2.6458 | -82.3267
2.6461 | -82.3214
2.6464 | -82.3163
2.6468 | -82.311
2.6471 | -82.3035
2.6474 | -82.2937
2.6477 | -82.2886
2.648 | -82.2846
2.6483 | -82.2796
2.6486 | -82.2754
2.6489 | -82.2727
2.6492 | -82.2691
2.6495 | -82.2649
2.6498 | -82.2606
2.6502 | -82.2552
2.6505 | -82.2504
2.6508 | -82.2457
2.6511 | -82.2415
2.6514 | -82.2387
2.6517 | -82.2359
2.652 | -82.2314
2.6523 | -82.2235
2.6526 | -82.2184
2.6529 | -82.2148
2.6532 | -82.211
2.6536 | -82.207
2.6539 | -82.2029
2.6542 | -82.1986
2.6545 | -82.1943
2.6548 | -82.1892
2.6551 | -82.184
2.6554 | -82.1795
2.6557 | -82.1748
2.656 | -82.1706
2.6563 | -82.1668
2.6566 | -82.1644
2.6569 | -82.1605
2.6573 | -82.1559
2.6576 | -82.1531
2.6579 | -82.1491
2.6582 | -82.1456
2.6585 | -82.1424
2.6588 | -82.1383
2.6591 | -82.1336
2.6594 | -82.1287
2.6597 | -82.1249
2.66 | -82.1217
2.6603 | -82.1183
2.6607 | -82.1143
2.661 | -82.1107
2.6613 | -82.1076
2.6616 | -82.103
2.6619 | -82.0963
2.6622 | -82.0895
2.6625 | -82.0834
2.6628 | -82.0791
2.6631 | -82.0734
2.6634 | -82.0678
2.6637 | -82.0616
2.664 | -82.0567
2.6644 | -82.0526
2.6647 | -82.0489
2.665 | -82.0446
2.6653 | -82.0409
2.6656 | -82.0369
2.6659 | -82.0332
2.6662 | -82.0254
2.6665 | -82.0167
2.6668 | -82.0123
2.6671 | -82.0072
2.6674 | -82.002
2.6678 | -81.9979
2.6681 | -81.9947
2.6684 | -81.9923
2.6687 | -81.9895
2.669 | -81.9858
2.6693 | -81.9808
2.6696 | -81.9765
2.6699 | -81.9722
2.6702 | -81.9676
2.6705 | -81.9633
2.6708 | -81.9589
2.6711 | -81.9525
2.6715 | -81.9472
2.6718 | -81.9436
2.6721 | -81.9388
2.6724 | -81.9329
2.6727 | -81.9287
2.673 | -81.9244
2.6733 | -81.9183
2.6736 | -81.9121
2.6739 | -81.9095
2.6742 | -81.9065
2.6745 | -81.9026
2.6749 | -81.8987
2.6752 | -81.8948
2.6755 | -81.8907
2.6758 | -81.8849
2.6761 | -81.8801
2.6764 | -81.8754
2.6767 | -81.8718
2.677 | -81.869
2.6773 | -81.867
2.6776 | -81.865
2.6779 | -81.8604
2.6783 | -81.8534
2.6786 | -81.8465
2.6789 | -81.8429
2.6792 | -81.8395
2.6795 | -81.8358
2.6798 | -81.834
2.6801 | -81.8306
2.6804 | -81.8261
2.6807 | -81.8211
2.681 | -81.8165
2.6813 | -81.814
2.6816 | -81.8091
2.682 | -81.8041
2.6823 | -81.8005
2.6826 | -81.797
2.6829 | -81.7919
2.6832 | -81.7881
2.6835 | -81.7842
2.6838 | -81.7793
2.6841 | -81.7756
2.6844 | -81.7729
2.6847 | -81.7683
2.685 | -81.764
2.6854 | -81.7615
2.6857 | -81.7588
2.686 | -81.7542
2.6863 | -81.7519
2.6866 | -81.7477
2.6869 | -81.7435
2.6872 | -81.7402
2.6875 | -81.7365
2.6878 | -81.7326
2.6881 | -81.729
2.6884 | -81.7257
2.6888 | -81.7217
2.6891 | -81.718
2.6894 | -81.7145
2.6897 | -81.7079
2.69 | -81.7033
2.6903 | -81.7001
2.6906 | -81.6965
2.6909 | -81.6926
2.6912 | -81.6889
2.6915 | -81.6851
2.6918 | -81.6818
2.6921 | -81.6782
2.6925 | -81.6751
2.6928 | -81.6724
2.6931 | -81.6695
2.6934 | -81.6664
2.6937 | -81.6628
2.694 | -81.6607
2.6943 | -81.6582
2.6946 | -81.6547
2.6949 | -81.6512
2.6952 | -81.646
2.6955 | -81.6418
2.6959 | -81.6361
2.6962 | -81.631
2.6965 | -81.6267
2.6968 | -81.6237
2.6971 | -81.6197
2.6974 | -81.6165
2.6977 | -81.6134
2.698 | -81.6109
2.6983 | -81.607
2.6986 | -81.6041
2.6989 | -81.5983
2.6992 | -81.5945
2.6996 | -81.5906
2.6999 | -81.5851
2.7002 | -81.5812
2.7005 | -81.5774
2.7008 | -81.5737
2.7011 | -81.5723
2.7014 | -81.5697
2.7017 | -81.5652
2.702 | -81.5613
2.7023 | -81.5581
2.7026 | -81.555
2.703 | -81.5515
2.7033 | -81.5475
2.7036 | -81.5437
2.7039 | -81.5389
2.7042 | -81.5341
2.7045 | -81.5311
2.7048 | -81.5275
2.7051 | -81.5245
2.7054 | -81.5206
2.7057 | -81.5185
2.706 | -81.5152
2.7064 | -81.5105
2.7067 | -81.5067
2.707 | -81.5042
2.7073 | -81.5022
2.7076 | -81.4991
2.7079 | -81.497
2.7082 | -81.4948
2.7085 | -81.4896
2.7088 | -81.485
2.7091 | -81.4822
2.7094 | -81.4793
2.7097 | -81.4764
2.7101 | -81.4713
2.7104 | -81.4684
2.7107 | -81.4652
2.711 | -81.4627
2.7113 | -81.4597
2.7116 | -81.4555
2.7119 | -81.4525
2.7122 | -81.4502
2.7125 | -81.4472
2.7128 | -81.4447
2.7131 | -81.4421
2.7135 | -81.4401
2.7138 | -81.4375
2.7141 | -81.4315
2.7144 | -81.4279
2.7147 | -81.425
2.715 | -81.4224
2.7153 | -81.419
2.7156 | -81.4156
2.7159 | -81.4122
2.7162 | -81.4104
2.7165 | -81.4096
2.7168 | -81.4088
2.7172 | -81.4071
2.7175 | -81.4046
2.7178 | -81.4019
2.7181 | -81.401
2.7184 | -81.3995
2.7187 | -81.3969
2.719 | -81.3943
2.7193 | -81.3932
2.7196 | -81.3904
2.7199 | -81.3863
2.7202 | -81.3838
2.7206 | -81.3812
2.7209 | -81.3796
2.7212 | -81.3774
2.7215 | -81.3749
2.7218 | -81.3728
2.7221 | -81.3704
2.7224 | -81.3669
2.7227 | -81.3634
2.723 | -81.3612
2.7233 | -81.3591
2.7236 | -81.3569
2.7239 | -81.3528
2.7243 | -81.3502
2.7246 | -81.3484
2.7249 | -81.3451
2.7252 | -81.3409
2.7255 | -81.3379
2.7258 | -81.337
2.7261 | -81.3349
2.7264 | -81.3307
2.7267 | -81.3275
2.727 | -81.3256
2.7273 | -81.3234
2.7277 | -81.3215
2.728 | -81.3204
2.7283 | -81.3186
2.7286 | -81.3166
2.7289 | -81.3143
2.7292 | -81.3132
2.7295 | -81.3127
2.7298 | -81.3103
2.7301 | -81.3085
2.7304 | -81.3058
2.7307 | -81.3039
2.7311 | -81.3022
2.7314 | -81.3006
2.7317 | -81.2992
2.732 | -81.2976
2.7323 | -81.2951
2.7326 | -81.2939
2.7329 | -81.2918
2.7332 | -81.2905
2.7335 | -81.2888
2.7338 | -81.2869
2.7341 | -81.2864
2.7344 | -81.2852
2.7348 | -81.2854
2.7351 | -81.2836
2.7354 | -81.28
2.7357 | -81.2784
2.736 | -81.2771
2.7363 | -81.2754
2.7366 | -81.2735
2.7369 | -81.2733
2.7372 | -81.2728
2.7375 | -81.2729
2.7378 | -81.2708
2.7382 | -81.2683
2.7385 | -81.2684
2.7388 | -81.2682
2.7391 | -81.267
2.7394 | -81.2654
2.7397 | -81.2648
2.74 | -81.2639
2.7403 | -81.2626
2.7406 | -81.2619
2.7409 | -81.2626
2.7412 | -81.2631
2.7415 | -81.262
2.7419 | -81.2614
2.7422 | -81.2609
2.7425 | -81.2605
2.7428 | -81.2604
2.7431 | -81.261
2.7434 | -81.2612
2.7437 | -81.2624
2.744 | -81.2621
2.7443 | -81.2612
2.7446 | -81.2607
2.7449 | -81.2614
2.7453 | -81.2631
2.7456 | -81.2629
2.7459 | -81.2621
2.7462 | -81.2629
2.7465 | -81.2624
2.7468 | -81.2623
2.7471 | -81.2624
2.7474 | -81.2617
2.7477 | -81.2621
2.748 | -81.2637
2.7483 | -81.2661
2.7487 | -81.266
2.749 | -81.2655
2.7493 | -81.265
2.7496 | -81.2632
2.7499 | -81.263
2.7502 | -81.2627
2.7505 | -81.2625
2.7508 | -81.2635
2.7511 | -81.2642
2.7514 | -81.2637
2.7517 | -81.2656
2.752 | -81.2685
2.7524 | -81.2659
2.7527 | -81.2654
2.753 | -81.2657
2.7533 | -81.2663
2.7536 | -81.2639
2.7539 | -81.2625
2.7542 | -81.2632
2.7545 | -81.2632
2.7548 | -81.2636
2.7551 | -81.2645
2.7554 | -81.2642
2.7558 | -81.2644
2.7561 | -81.2646
2.7564 | -81.2627
2.7567 | -81.2647
2.757 | -81.2656
2.7573 | -81.2669
2.7576 | -81.267
2.7579 | -81.2666
2.7582 | -81.2668
2.7585 | -81.2675
2.7588 | -81.2689
2.7591 | -81.2701
2.7595 | -81.273
2.7598 | -81.2735
2.7601 | -81.2741
2.7604 | -81.2742
2.7607 | -81.2739
2.761 | -81.2746
2.7613 | -81.2754
2.7616 | -81.2781
2.7619 | -81.2793
2.7622 | -81.2808
2.7625 | -81.2834
2.7629 | -81.2856
2.7632 | -81.2854
2.7635 | -81.2857
2.7638 | -81.2862
2.7641 | -81.2872
2.7644 | -81.2887
2.7647 | -81.2899
2.765 | -81.2909
2.7653 | -81.2922
2.7656 | -81.2924
2.7659 | -81.2926
2.7662 | -81.294
2.7666 | -81.2936
2.7669 | -81.2944
2.7672 | -81.296
2.7675 | -81.2983
2.7678 | -81.3
2.7681 | -81.3012
2.7684 | -81.3018
2.7687 | -81.3027
2.769 | -81.303
2.7693 | -81.3043
2.7696 | -81.3064
2.77 | -81.3078
2.7703 | -81.3094
2.7706 | -81.3092
2.7709 | -81.3112
2.7712 | -81.3136
2.7715 | -81.3164
2.7718 | -81.3182
2.7721 | -81.3181
2.7724 | -81.3182
2.7727 | -81.3184
2.773 | -81.3198
2.7734 | -81.3222
2.7737 | -81.3245
2.774 | -81.3262
2.7743 | -81.3285
2.7746 | -81.3302
2.7749 | -81.3319
2.7752 | -81.3338
2.7755 | -81.3354
2.7758 | -81.3376
2.7761 | -81.3395
2.7764 | -81.341
2.7767 | -81.3421
2.7771 | -81.3435
2.7774 | -81.3453
2.7777 | -81.3471
2.778 | -81.3495
2.7783 | -81.351
2.7786 | -81.3519
2.7789 | -81.3527
2.7792 | -81.3534
2.7795 | -81.3533
2.7798 | -81.3552
2.7801 | -81.356
2.7805 | -81.3576
2.7808 | -81.3597
2.7811 | -81.3618
2.7814 | -81.3631
2.7817 | -81.3658
2.782 | -81.3676
2.7823 | -81.3676
2.7826 | -81.3684
2.7829 | -81.3707
2.7832 | -81.3735
2.7835 | -81.3755
2.7839 | -81.3788
2.7842 | -81.3808
2.7845 | -81.3807
2.7848 | -81.3822
2.7851 | -81.3842
2.7854 | -81.3845
2.7857 | -81.3858
2.786 | -81.3884
2.7863 | -81.3898
2.7866 | -81.3915
2.7869 | -81.3931
2.7872 | -81.3947
2.7876 | -81.3975
2.7879 | -81.4004
2.7882 | -81.4029
2.7885 | -81.405
2.7888 | -81.4073
2.7891 | -81.41
2.7894 | -81.4137
2.7897 | -81.4171
2.79 | -81.4187
2.7903 | -81.4195
2.7906 | -81.4202
2.791 | -81.422
2.7913 | -81.4253
2.7916 | -81.4256
2.7919 | -81.4202
2.7922 | -81.4176
2.7925 | -81.42
2.7928 | -81.4213
2.7931 | -81.4212
2.7934 | -81.4208
2.7937 | -81.4225
2.794 | -81.4238
2.7943 | -81.4248
2.7947 | -81.4248
2.795 | -81.4241
2.7953 | -81.4253
2.7956 | -81.4258
2.7959 | -81.4273
2.7962 | -81.4289
2.7965 | -81.4294
2.7968 | -81.4284
2.7971 | -81.429
2.7974 | -81.4294
2.7977 | -81.4303
2.7981 | -81.4303
2.7984 | -81.4318
2.7987 | -81.4327
2.799 | -81.4335
2.7993 | -81.4342
2.7996 | -81.4345
2.7999 | -81.436
2.8002 | -81.436
2.8005 | -81.436
2.8008 | -81.4369
2.8011 | -81.4378
2.8014 | -81.4377
2.8018 | -81.4375
2.8021 | -81.4376
2.8024 | -81.4378
2.8027 | -81.4381
2.803 | -81.4395
2.8033 | -81.4393
2.8036 | -81.4407
2.8039 | -81.4424
2.8042 | -81.4429
2.8045 | -81.4431
2.8048 | -81.4428
2.8052 | -81.4438
2.8055 | -81.4446
2.8058 | -81.4458
2.8061 | -81.4449
2.8064 | -81.4473
2.8067 | -81.4476
2.807 | -81.4485
2.8073 | -81.4496
2.8076 | -81.4495
2.8079 | -81.45
2.8082 | -81.4492
2.8086 | -81.4494
2.8089 | -81.4492
2.8092 | -81.449
2.8095 | -81.4491
2.8098 | -81.4494
2.8101 | -81.4498
2.8104 | -81.4502
2.8107 | -81.4484
2.811 | -81.445
2.8113 | -81.4448
2.8116 | -81.445
2.8119 | -81.4448
2.8123 | -81.4434
2.8126 | -81.4418
2.8129 | -81.4431
2.8132 | -81.4439
2.8135 | -81.4424
2.8138 | -81.4425
2.8141 | -81.4435
2.8144 | -81.4426
2.8147 | -81.4413
2.815 | -81.4401
2.8153 | -81.4397
2.8157 | -81.4363
2.816 | -81.4358
2.8163 | -81.4373
2.8166 | -81.4374
2.8169 | -81.4375
2.8172 | -81.4367
2.8175 | -81.4339
2.8178 | -81.43
2.8181 | -81.4293
2.8184 | -81.427
2.8187 | -81.4263
2.819 | -81.4258
2.8194 | -81.4265
2.8197 | -81.4258
2.82 | -81.4247
2.8203 | -81.4241
2.8206 | -81.4214
2.8209 | -81.4201
2.8212 | -81.4197
2.8215 | -81.4182
2.8218 | -81.417
2.8221 | -81.4161
2.8224 | -81.4134
2.8228 | -81.413
2.8231 | -81.4126
2.8234 | -81.4109
2.8237 | -81.409
2.824 | -81.4055
2.8243 | -81.4033
2.8246 | -81.4009
2.8249 | -81.3993
2.8252 | -81.3971
2.8255 | -81.3966
2.8258 | -81.3948
2.8262 | -81.3911
2.8265 | -81.3828
2.8268 | -81.3786
2.8271 | -81.3766
2.8274 | -81.3742
2.8277 | -81.3719
2.828 | -81.3702
2.8283 | -81.3681
2.8286 | -81.3656
2.8289 | -81.3632
2.8292 | -81.3612
2.8295 | -81.3597
2.8299 | -81.3572
2.8302 | -81.3545
2.8305 | -81.3527
2.8308 | -81.35
2.8311 | -81.3466
2.8314 | -81.3468
2.8317 | -81.345
2.832 | -81.3431
2.8323 | -81.3407
2.8326 | -81.3402
2.8329 | -81.3368
2.8333 | -81.3351
2.8336 | -81.3328
2.8339 | -81.3303
2.8342 | -81.3271
2.8345 | -81.3249
2.8348 | -81.3215
2.8351 | -81.314
2.8354 | -81.3098
2.8357 | -81.3068
2.836 | -81.3025
2.8363 | -81.299
2.8366 | -81.2977
2.837 | -81.2954
2.8373 | -81.2923
2.8376 | -81.2891
2.8379 | -81.2873
2.8382 | -81.284
2.8385 | -81.2745
2.8388 | -81.2695
2.8391 | -81.2664
2.8394 | -81.263
2.8397 | -81.2595
2.84 | -81.2556
2.8404 | -81.2513
2.8407 | -81.2469
2.841 | -81.2444
2.8413 | -81.2412
2.8416 | -81.2373
2.8419 | -81.2341
2.8422 | -81.2296
2.8425 | -81.2238
2.8428 | -81.2188
2.8431 | -81.2143
2.8434 | -81.2081
2.8438 | -81.2022
2.8441 | -81.1977
2.8444 | -81.1931
2.8447 | -81.1889
2.845 | -81.1855
2.8453 | -81.182
2.8456 | -81.1773
2.8459 | -81.1741
2.8462 | -81.1684
2.8465 | -81.1592
2.8468 | -81.1537
2.8471 | -81.1496
2.8475 | -81.145
2.8478 | -81.1401
2.8481 | -81.1354
2.8484 | -81.1299
2.8487 | -81.1256
2.849 | -81.1215
2.8493 | -81.117
2.8496 | -81.1123
2.8499 | -81.1081
2.8502 | -81.1033
2.8505 | -81.0973
2.8509 | -81.0932
2.8512 | -81.0882
2.8515 | -81.0819
2.8518 | -81.0772
2.8521 | -81.0709
2.8524 | -81.0622
2.8527 | -81.056
2.853 | -81.0511
2.8533 | -81.0469
2.8536 | -81.0435
2.8539 | -81.041
2.8542 | -81.037
2.8546 | -81.0328
2.8549 | -81.03
2.8552 | -81.0261
2.8555 | -81.0217
2.8558 | -81.0181
2.8561 | -81.013
2.8564 | -81.0084
2.8567 | -81.0035
2.857 | -80.9986
2.8573 | -80.9934
2.8576 | -80.9892
2.858 | -80.984
2.8583 | -80.9795
2.8586 | -80.9772
2.8589 | -80.9725
2.8592 | -80.9686
2.8595 | -80.964
2.8598 | -80.9598
2.8601 | -80.9551
2.8604 | -80.95
2.8607 | -80.945
2.861 | -80.94
2.8614 | -80.934
2.8617 | -80.9273
2.862 | -80.9202
2.8623 | -80.9152
2.8626 | -80.91
2.8629 | -80.9049
2.8632 | -80.9002
2.8635 | -80.8954
2.8638 | -80.8902
2.8641 | -80.885
2.8644 | -80.8807
2.8647 | -80.8766
2.8651 | -80.8706
2.8654 | -80.863
2.8657 | -80.8572
2.866 | -80.8541
2.8663 | -80.8505
2.8666 | -80.8447
2.8669 | -80.8391
2.8672 | -80.8322
2.8675 | -80.8208
2.8678 | -80.8124
2.8681 | -80.8076
2.8685 | -80.8028
2.8688 | -80.7993
2.8691 | -80.7936
2.8694 | -80.7865
2.8697 | -80.7811
2.87 | -80.7754
2.8703 | -80.7691
2.8706 | -80.7627
2.8709 | -80.7577
2.8712 | -80.7545
2.8715 | -80.7505
2.8718 | -80.7453
2.8722 | -80.7389
2.8725 | -80.7306
2.8728 | -80.7223
2.8731 | -80.7166
2.8734 | -80.7111
2.8737 | -80.7057
2.874 | -80.7002
2.8743 | -80.6947
2.8746 | -80.6892
2.8749 | -80.6819
2.8752 | -80.6751
2.8756 | -80.6696
2.8759 | -80.6648
2.8762 | -80.6602
2.8765 | -80.6537
2.8768 | -80.6477
2.8771 | -80.6429
2.8774 | -80.6387
2.8777 | -80.634
2.878 | -80.6287
2.8783 | -80.6238
2.8786 | -80.6187
2.879 | -80.6136
2.8793 | -80.6082
2.8796 | -80.6033
2.8799 | -80.5963
2.8802 | -80.5893
2.8805 | -80.5833
2.8808 | -80.5772
2.8811 | -80.5703
2.8814 | -80.5636
2.8817 | -80.5572
2.882 | -80.5515
2.8823 | -80.5452
2.8827 | -80.5377
2.883 | -80.5321
2.8833 | -80.5288
2.8836 | -80.5248
2.8839 | -80.5195
2.8842 | -80.5127
2.8845 | -80.5065
2.8848 | -80.4976
2.8851 | -80.4912
2.8854 | -80.4854
2.8857 | -80.4803
2.8861 | -80.4751
2.8864 | -80.4665
2.8867 | -80.4577
2.887 | -80.4513
2.8873 | -80.4449
2.8876 | -80.4398
2.8879 | -80.4345
2.8882 | -80.4292
2.8885 | -80.4239
2.8888 | -80.4169
2.8891 | -80.4116
2.8894 | -80.4072
2.8898 | -80.4019
2.8901 | -80.3982
2.8904 | -80.3954
2.8907 | -80.3917
2.891 | -80.3865
2.8913 | -80.3811
2.8916 | -80.3757
2.8919 | -80.3706
2.8922 | -80.3655
2.8925 | -80.3599
2.8928 | -80.354
2.8932 | -80.3498
2.8935 | -80.3458
2.8938 | -80.3407
2.8941 | -80.336
2.8944 | -80.3317
2.8947 | -80.326
2.895 | -80.3215
2.8953 | -80.3182
2.8956 | -80.3129
2.8959 | -80.3082
2.8962 | -80.304
2.8965 | -80.3
2.8969 | -80.2952
2.8972 | -80.2897
2.8975 | -80.2846
2.8978 | -80.2802
2.8981 | -80.2749
2.8984 | -80.2704
2.8987 | -80.2667
2.899 | -80.2615
2.8993 | -80.255
2.8996 | -80.2506
2.8999 | -80.2458
2.9003 | -80.242
2.9006 | -80.2371
2.9009 | -80.2271
2.9012 | -80.2205
2.9015 | -80.2165
2.9018 | -80.2108
2.9021 | -80.2039
2.9024 | -80.1978
2.9027 | -80.1918
2.903 | -80.1867
2.9033 | -80.1828
2.9037 | -80.1804
2.904 | -80.1782
2.9043 | -80.175
2.9046 | -80.1714
2.9049 | -80.1672
2.9052 | -80.1634
2.9055 | -80.1607
2.9058 | -80.1563
2.9061 | -80.1515
2.9064 | -80.1472
2.9067 | -80.1433
2.907 | -80.1395
2.9074 | -80.1353
2.9077 | -80.1316
2.908 | -80.1282
2.9083 | -80.1243
2.9086 | -80.1196
2.9089 | -80.1147
2.9092 | -80.1102
2.9095 | -80.1077
2.9098 | -80.1037
2.9101 | -80.0996
2.9104 | -80.0946
2.9108 | -80.0908
2.9111 | -80.0871
2.9114 | -80.0841
2.9117 | -80.0804
2.912 | -80.0765
2.9123 | -80.0731
2.9126 | -80.0696
2.9129 | -80.0674
2.9132 | -80.0638
2.9135 | -80.06
2.9138 | -80.0556
2.9141 | -80.0519
2.9145 | -80.0484
2.9148 | -80.0443
2.9151 | -80.0408
2.9154 | -80.038
2.9157 | -80.0352
2.916 | -80.0331
2.9163 | -80.029
2.9166 | -80.0253
2.9169 | -80.0212
2.9172 | -80.0174
2.9175 | -80.0145
2.9179 | -80.0114
2.9182 | -80.0077
2.9185 | -80.0047
2.9188 | -80.0012
2.9191 | -79.9982
2.9194 | -79.9955
2.9197 | -79.9928
2.92 | -79.991
2.9203 | -79.988
2.9206 | -79.9856
2.9209 | -79.9844
2.9213 | -79.983
2.9216 | -79.9805
2.9219 | -79.9794
2.9222 | -79.9765
2.9225 | -79.9746
2.9228 | -79.9726
2.9231 | -79.9718
2.9234 | -79.9698
2.9237 | -79.9672
2.924 | -79.9617
2.9243 | -79.9584
2.9246 | -79.9557
2.925 | -79.9536
2.9253 | -79.9506
2.9256 | -79.9484
2.9259 | -79.9466
2.9262 | -79.9444
2.9265 | -79.9428
2.9268 | -79.9409
2.9271 | -79.9395
2.9274 | -79.9379
2.9277 | -79.9356
2.928 | -79.9324
2.9284 | -79.9289
2.9287 | -79.9258
2.929 | -79.9235
2.9293 | -79.921
2.9296 | -79.92
2.9299 | -79.9187
2.9302 | -79.9169
2.9305 | -79.9143
2.9308 | -79.9135
2.9311 | -79.9117
2.9314 | -79.9105
2.9318 | -79.9093
2.9321 | -79.9079
2.9324 | -79.9068
2.9327 | -79.9068
2.933 | -79.905
2.9333 | -79.9022
2.9336 | -79.8987
2.9339 | -79.8969
2.9342 | -79.8945
2.9345 | -79.8923
2.9348 | -79.8909
2.9351 | -79.8903
2.9355 | -79.8896
2.9358 | -79.8874
2.9361 | -79.8857
2.9364 | -79.8849
2.9367 | -79.8839
2.937 | -79.8828
2.9373 | -79.8816
2.9376 | -79.8805
2.9379 | -79.88
2.9382 | -79.8787
2.9385 | -79.876
2.9389 | -79.8743
2.9392 | -79.8731
2.9395 | -79.8718
2.9398 | -79.8734
2.9401 | -79.8716
2.9404 | -79.8706
2.9407 | -79.8685
2.941 | -79.8674
2.9413 | -79.867
2.9416 | -79.8668
2.9419 | -79.8664
2.9422 | -79.8665
2.9426 | -79.8664
2.9429 | -79.8661
2.9432 | -79.8662
2.9435 | -79.8657
2.9438 | -79.8654
2.9441 | -79.8643
2.9444 | -79.8636
2.9447 | -79.8621
2.945 | -79.8613
2.9453 | -79.8616
2.9456 | -79.8615
2.946 | -79.8614
2.9463 | -79.8615
2.9466 | -79.8622
2.9469 | -79.8622
2.9472 | -79.8635
2.9475 | -79.8651
2.9478 | -79.8652
2.9481 | -79.8641
2.9484 | -79.862
2.9487 | -79.8633
2.949 | -79.8634
2.9493 | -79.8638
2.9497 | -79.8646
2.95 | -79.865
2.9503 | -79.8655
2.9506 | -79.8683
2.9509 | -79.8709
2.9512 | -79.8705
2.9515 | -79.8712
2.9518 | -79.8715
2.9521 | -79.8739
2.9524 | -79.8761
2.9527 | -79.8759
2.9531 | -79.8764
2.9534 | -79.8768
2.9537 | -79.8769
2.954 | -79.8783
2.9543 | -79.8789
2.9546 | -79.88
2.9549 | -79.8824
2.9552 | -79.8828
2.9555 | -79.8839
2.9558 | -79.8854
2.9561 | -79.8872
2.9565 | -79.8889
2.9568 | -79.8899
2.9571 | -79.8908
2.9574 | -79.8912
2.9577 | -79.8942
2.958 | -79.8938
2.9583 | -79.8939
2.9586 | -79.8947
2.9589 | -79.8949
2.9592 | -79.8966
2.9595 | -79.8977
2.9598 | -79.8974
2.9602 | -79.898
2.9605 | -79.8992
2.9608 | -79.9003
2.9611 | -79.9006
2.9614 | -79.8994
2.9617 | -79.8993
2.962 | -79.9008
2.9623 | -79.9024
2.9626 | -79.9042
2.9629 | -79.9057
2.9632 | -79.907
2.9636 | -79.9085
2.9639 | -79.9105
2.9642 | -79.9123
2.9645 | -79.9139
2.9648 | -79.9149
2.9651 | -79.9162
2.9654 | -79.9174
2.9657 | -79.9196
2.966 | -79.9216
2.9663 | -79.9249
2.9666 | -79.9265
2.9669 | -79.928
2.9673 | -79.9303
2.9676 | -79.9313
2.9679 | -79.9326
2.9682 | -79.9354
2.9685 | -79.9386
2.9688 | -79.9414
2.9691 | -79.9416
2.9694 | -79.9416
2.9697 | -79.9435
2.97 | -79.9467
2.9703 | -79.9499
2.9707 | -79.9519
2.971 | -79.9529
2.9713 | -79.9525
2.9716 | -79.9541
2.9719 | -79.9572
2.9722 | -79.9604
2.9725 | -79.9632
2.9728 | -79.9655
2.9731 | -79.9665
2.9734 | -79.968
2.9737 | -79.9708
2.9741 | -79.9729
2.9744 | -79.9751
2.9747 | -79.9762
2.975 | -79.9769
2.9753 | -79.9809
2.9756 | -79.9832
2.9759 | -79.9858
2.9762 | -79.9884
2.9765 | -79.9907
2.9768 | -79.9929
2.9771 | -79.995
2.9774 | -79.9975
2.9778 | -79.9998
2.9781 | -80.0024
2.9784 | -80.0048
2.9787 | -80.0076
2.979 | -80.011
2.9793 | -80.0134
2.9796 | -80.0156
2.9799 | -80.0181
2.9802 | -80.0207
2.9805 | -80.024
2.9808 | -80.0257
2.9812 | -80.0273
2.9815 | -80.0296
2.9818 | -80.0324
2.9821 | -80.0352
2.9824 | -80.0379
2.9827 | -80.0403
2.983 | -80.0428
2.9833 | -80.0444
2.9836 | -80.0463
2.9839 | -80.05
2.9842 | -80.0537
2.9845 | -80.0558
2.9849 | -80.0575
2.9852 | -80.0612
2.9855 | -80.064
2.9858 | -80.0668
2.9861 | -80.0704
2.9864 | -80.0735
2.9867 | -80.0762
2.987 | -80.0791
2.9873 | -80.0813
2.9876 | -80.0822
2.9879 | -80.0848
2.9883 | -80.0869
2.9886 | -80.0887
2.9889 | -80.0918
2.9892 | -80.0942
2.9895 | -80.097
2.9898 | -80.0998
2.9901 | -80.1025
2.9904 | -80.1045
2.9907 | -80.108
2.991 | -80.1118
2.9913 | -80.1154
2.9917 | -80.1187
2.992 | -80.1215
2.9923 | -80.1246
2.9926 | -80.1289
2.9929 | -80.137
2.9932 | -80.141
2.9935 | -80.1469
2.9938 | -80.15
2.9941 | -80.1536
2.9944 | -80.1574
2.9947 | -80.1623
2.995 | -80.166
2.9954 | -80.169
2.9957 | -80.1723
2.996 | -80.1761
2.9963 | -80.1808
2.9966 | -80.184
2.9969 | -80.188
2.9972 | -80.1911
2.9975 | -80.1944
2.9978 | -80.1979
2.9981 | -80.2026
2.9984 | -80.2041
2.9988 | -80.2055
2.9991 | -80.209
2.9994 | -80.2133
2.9997 | -80.2175
3 | -80.2208
3.0003 | -80.2243
3.0006 | -80.2273
3.0009 | -80.2299
3.0012 | -80.2339
3.0015 | -80.2367
3.0018 | -80.2403
3.0021 | -80.2444
3.0025 | -80.2483
3.0028 | -80.2517
3.0031 | -80.255
3.0034 | -80.2585
3.0037 | -80.2603
3.004 | -80.2637
3.0043 | -80.267
3.0046 | -80.2702
3.0049 | -80.2728
3.0052 | -80.2766
3.0055 | -80.2802
3.0059 | -80.2835
3.0062 | -80.2871
3.0065 | -80.291
3.0068 | -80.2943
3.0071 | -80.296
3.0074 | -80.2995
3.0077 | -80.3035
3.008 | -80.3067
3.0083 | -80.3104
3.0086 | -80.3128
3.0089 | -80.3156
3.0093 | -80.3183
3.0096 | -80.3221
3.0099 | -80.3249
3.0102 | -80.3281
3.0105 | -80.332
3.0108 | -80.3354
3.0111 | -80.3398
3.0114 | -80.3436
3.0117 | -80.3474
3.012 | -80.3512
3.0123 | -80.3537
3.0126 | -80.3568
3.013 | -80.36
3.0133 | -80.3649
3.0136 | -80.3707
3.0139 | -80.3745
3.0142 | -80.3775
3.0145 | -80.3813
3.0148 | -80.3857
3.0151 | -80.3896
3.0154 | -80.3934
3.0157 | -80.3974
3.016 | -80.4022
3.0164 | -80.4061
3.0167 | -80.4093
3.017 | -80.4145
3.0173 | -80.419
3.0176 | -80.4204
3.0179 | -80.4232
3.0182 | -80.4265
3.0185 | -80.4319
3.0188 | -80.437
3.0191 | -80.4408
3.0194 | -80.4452
3.0197 | -80.4497
3.0201 | -80.4528
3.0204 | -80.456
3.0207 | -80.4609
3.021 | -80.4666
3.0213 | -80.4701
3.0216 | -80.4751
3.0219 | -80.4786
3.0222 | -80.4821
3.0225 | -80.4845
3.0228 | -80.488
3.0231 | -80.4924
3.0235 | -80.4953
3.0238 | -80.4998
3.0241 | -80.5031
3.0244 | -80.5064
3.0247 | -80.5098
3.025 | -80.5137
3.0253 | -80.5189
3.0256 | -80.5251
3.0259 | -80.5279
3.0262 | -80.5312
3.0265 | -80.5347
3.0269 | -80.5396
3.0272 | -80.5424
3.0275 | -80.5474
3.0278 | -80.5513
3.0281 | -80.5557
3.0284 | -80.5595
3.0287 | -80.5639
3.029 | -80.5685
3.0293 | -80.5722
3.0296 | -80.576
3.0299 | -80.5818
3.0302 | -80.5909
3.0306 | -80.5951
3.0309 | -80.5981
3.0312 | -80.6014
3.0315 | -80.603
3.0318 | -80.6038
3.0321 | -80.6073
3.0324 | -80.6108
3.0327 | -80.6147
3.033 | -80.617
3.0333 | -80.6204
3.0336 | -80.6245
3.034 | -80.6286
3.0343 | -80.6322
3.0346 | -80.6357
3.0349 | -80.6412
3.0352 | -80.6451
3.0355 | -80.6497
3.0358 | -80.6524
3.0361 | -80.6532
3.0364 | -80.6564
3.0367 | -80.6598
3.037 | -80.6639
3.0373 | -80.6674
3.0377 | -80.6718
3.038 | -80.6752
3.0383 | -80.6787
3.0386 | -80.6833
3.0389 | -80.6864
3.0392 | -80.689
3.0395 | -80.6971
3.0398 | -80.7018
3.0401 | -80.7053
3.0404 | -80.7097
3.0407 | -80.7144
3.0411 | -80.7181
3.0414 | -80.7216
3.0417 | -80.725
3.042 | -80.7288
3.0423 | -80.7309
3.0426 | -80.7348
3.0429 | -80.7438
3.0432 | -80.7469
3.0435 | -80.7498
3.0438 | -80.7529
3.0441 | -80.7562
3.0444 | -80.76
3.0448 | -80.7637
3.0451 | -80.7673
3.0454 | -80.77
3.0457 | -80.7739
3.046 | -80.7767
3.0463 | -80.7782
3.0466 | -80.7807
3.0469 | -80.7828
3.0472 | -80.7869
3.0475 | -80.7922
3.0478 | -80.7963
3.0482 | -80.7992
3.0485 | -80.8066
3.0488 | -80.8126
3.0491 | -80.8164
3.0494 | -80.8203
3.0497 | -80.8239
3.05 | -80.8274
3.0503 | -80.8312
3.0506 | -80.8353
3.0509 | -80.8386
3.0512 | -80.8446
3.0516 | -80.853
3.0519 | -80.8586
3.0522 | -80.8627
3.0525 | -80.8664
3.0528 | -80.8699
3.0531 | -80.8733
3.0534 | -80.8784
3.0537 | -80.8816
3.054 | -80.8847
3.0543 | -80.8883
3.0546 | -80.8919
3.0549 | -80.895
3.0553 | -80.8975
3.0556 | -80.9015
3.0559 | -80.9059
3.0562 | -80.91
3.0565 | -80.9139
3.0568 | -80.9177
3.0571 | -80.922
3.0574 | -80.9251
3.0577 | -80.9302
3.058 | -80.9378
3.0583 | -80.9427
3.0587 | -80.9455
3.059 | -80.9477
3.0593 | -80.9513
3.0596 | -80.9526
3.0599 | -80.9535
3.0602 | -80.9543
3.0605 | -80.9571
3.0608 | -80.9611
3.0611 | -80.965
3.0614 | -80.9692
3.0617 | -80.9738
3.0621 | -80.9769
3.0624 | -80.9801
3.0627 | -80.984
3.063 | -80.9879
3.0633 | -80.9911
3.0636 | -80.9939
3.0639 | -81.0001
3.0642 | -81.0047
3.0645 | -81.0095
3.0648 | -81.0133
3.0651 | -81.0171
3.0654 | -81.0208
3.0658 | -81.024
3.0661 | -81.0267
3.0664 | -81.0288
3.0667 | -81.0307
3.067 | -81.0332
3.0673 | -81.036
3.0676 | -81.0403
3.0679 | -81.0464
3.0682 | -81.0516
3.0685 | -81.0547
3.0688 | -81.0585
3.0692 | -81.0619
3.0695 | -81.0639
3.0698 | -81.0688
3.0701 | -81.0728
3.0704 | -81.0761
3.0707 | -81.0794
3.071 | -81.0833
3.0713 | -81.0884
3.0716 | -81.092
3.0719 | -81.0933
3.0722 | -81.0965
3.0725 | -81.0997
3.0729 | -81.1027
3.0732 | -81.105
3.0735 | -81.1082
3.0738 | -81.1128
3.0741 | -81.1215
3.0744 | -81.1276
3.0747 | -81.1313
3.075 | -81.1336
3.0753 | -81.1356
3.0756 | -81.1378
3.0759 | -81.1421
3.0763 | -81.1471
3.0766 | -81.1504
3.0769 | -81.1536
3.0772 | -81.1574
3.0775 | -81.1617
3.0778 | -81.1651
3.0781 | -81.1678
3.0784 | -81.1709
3.0787 | -81.1741
3.079 | -81.1782
3.0793 | -81.1844
3.0796 | -81.1902
3.08 | -81.1934
3.0803 | -81.1975
3.0806 | -81.2018
3.0809 | -81.2056
3.0812 | -81.2087
3.0815 | -81.2118
3.0818 | -81.2171
3.0821 | -81.2208
3.0824 | -81.2266
3.0827 | -81.23
3.083 | -81.2333
3.0834 | -81.2369
3.0837 | -81.241
3.084 | -81.244
3.0843 | -81.2474
3.0846 | -81.2503
3.0849 | -81.2536
3.0852 | -81.2565
3.0855 | -81.2599
3.0858 | -81.263
3.0861 | -81.2665
3.0864 | -81.2696
3.0868 | -81.2729
3.0871 | -81.2764
3.0874 | -81.2812
3.0877 | -81.287
3.088 | -81.2884
3.0883 | -81.2891
3.0886 | -81.2939
3.0889 | -81.2983
3.0892 | -81.3027
3.0895 | -81.3054
3.0898 | -81.3089
3.0901 | -81.3139
3.0905 | -81.3176
3.0908 | -81.3208
3.0911 | -81.3245
3.0914 | -81.327
3.0917 | -81.3298
3.092 | -81.3341
3.0923 | -81.3384
3.0926 | -81.3461
3.0929 | -81.3497
3.0932 | -81.3529
3.0935 | -81.3554
3.0939 | -81.358
3.0942 | -81.3615
3.0945 | -81.3693
3.0948 | -81.3751
3.0951 | -81.3791
3.0954 | -81.3837
3.0957 | -81.3864
3.096 | -81.3897
3.0963 | -81.3926
3.0966 | -81.3961
3.0969 | -81.4007
3.0972 | -81.4033
3.0976 | -81.4066
3.0979 | -81.4102
3.0982 | -81.4132
3.0985 | -81.4154
3.0988 | -81.4163
3.0991 | -81.4169
3.0994 | -81.4174
3.0997 | -81.4194
3.1 | -81.4209
3.1003 | -81.4237
3.1006 | -81.4261
3.101 | -81.4292
3.1013 | -81.4328
3.1016 | -81.4364
3.1019 | -81.4384
3.1022 | -81.4408
3.1025 | -81.4431
3.1028 | -81.4461
3.1031 | -81.4491
3.1034 | -81.4493
3.1037 | -81.4514
3.104 | -81.4542
3.1044 | -81.4569
3.1047 | -81.4577
3.105 | -81.4588
3.1053 | -81.4609
3.1056 | -81.464
3.1059 | -81.4669
3.1062 | -81.4695
3.1065 | -81.472
3.1068 | -81.4758
3.1071 | -81.4786
3.1074 | -81.4814
3.1077 | -81.4859
3.1081 | -81.4904
3.1084 | -81.4917
3.1087 | -81.494
3.109 | -81.496
3.1093 | -81.4975
3.1096 | -81.5008
3.1099 | -81.5095
3.1102 | -81.515
3.1105 | -81.5172
3.1108 | -81.5209
3.1111 | -81.5264
3.1115 | -81.529
3.1118 | -81.5326
3.1121 | -81.5352
3.1124 | -81.5371
3.1127 | -81.5388
3.113 | -81.5393
3.1133 | -81.5404
3.1136 | -81.5421
3.1139 | -81.5447
3.1142 | -81.5468
3.1145 | -81.5488
3.1148 | -81.551
3.1152 | -81.5522
3.1155 | -81.5541
3.1158 | -81.5557
3.1161 | -81.5561
3.1164 | -81.5568
3.1167 | -81.5588
3.117 | -81.5609
3.1173 | -81.5626
3.1176 | -81.5628
3.1179 | -81.5646
3.1182 | -81.5672
3.1186 | -81.5709
3.1189 | -81.5741
3.1192 | -81.576
3.1195 | -81.5775
3.1198 | -81.579
3.1201 | -81.5818
3.1204 | -81.5838
3.1207 | -81.5837
3.121 | -81.5834
3.1213 | -81.5847
3.1216 | -81.5864
3.122 | -81.588
3.1223 | -81.5886
3.1226 | -81.5888
3.1229 | -81.5894
3.1232 | -81.591
3.1235 | -81.5923
3.1238 | -81.595
3.1241 | -81.5972
3.1244 | -81.5985
3.1247 | -81.6003
3.125 | -81.6023
3.1253 | -81.6034
3.1257 | -81.6046
3.126 | -81.6063
3.1263 | -81.6083
3.1266 | -81.6105
3.1269 | -81.6102
3.1272 | -81.6131
3.1275 | -81.6156
3.1278 | -81.6168
3.1281 | -81.6158
3.1284 | -81.6131
3.1287 | -81.6141
3.1291 | -81.6168
3.1294 | -81.6178
3.1297 | -81.6165
3.13 | -81.6169
3.1303 | -81.6174
3.1306 | -81.6159
3.1309 | -81.6174
3.1312 | -81.6165
3.1315 | -81.6152
3.1318 | -81.6154
3.1321 | -81.6177
3.1324 | -81.6187
3.1328 | -81.6199
3.1331 | -81.6207
3.1334 | -81.6205
3.1337 | -81.6209
3.134 | -81.6214
3.1343 | -81.622
3.1346 | -81.6272
3.1349 | -81.6297
3.1352 | -81.6303
3.1355 | -81.6301
3.1358 | -81.6299
3.1362 | -81.6318
3.1365 | -81.6326
3.1368 | -81.6321
3.1371 | -81.6325
3.1374 | -81.6307
3.1377 | -81.6294
3.138 | -81.6292
3.1383 | -81.6289
3.1386 | -81.6286
3.1389 | -81.6303
3.1392 | -81.633
3.1396 | -81.6322
3.1399 | -81.6317
3.1402 | -81.6326
3.1405 | -81.6314
3.1408 | -81.6281
3.1411 | -81.6218
3.1414 | -81.6191
3.1417 | -81.6213
3.142 | -81.6195
3.1423 | -81.6173
3.1426 | -81.6138
3.1429 | -81.6094
3.1433 | -81.6054
3.1436 | -81.6037
3.1439 | -81.6034
3.1442 | -81.6026
3.1445 | -81.6006
3.1448 | -81.5956
3.1451 | -81.5918
3.1454 | -81.5901
3.1457 | -81.5903
3.146 | -81.5918
3.1463 | -81.5898
3.1467 | -81.5841
3.147 | -81.5807
3.1473 | -81.579
3.1476 | -81.5758
3.1479 | -81.5712
3.1482 | -81.568
3.1485 | -81.565
3.1488 | -81.5622
3.1491 | -81.5609
3.1494 | -81.5558
3.1497 | -81.5513
3.15 | -81.552
3.1504 | -81.5477
3.1507 | -81.5433
3.151 | -81.5397
3.1513 | -81.5359
3.1516 | -81.5314
3.1519 | -81.5261
3.1522 | -81.5158
3.1525 | -81.5102
3.1528 | -81.5082
3.1531 | -81.5023
3.1534 | -81.496
3.1538 | -81.4929
3.1541 | -81.4889
3.1544 | -81.4857
3.1547 | -81.4823
3.155 | -81.4761
3.1553 | -81.4646
3.1556 | -81.4559
3.1559 | -81.4523
3.1562 | -81.4453
3.1565 | -81.44
3.1568 | -81.4382
3.1572 | -81.4265
3.1575 | -81.4122
3.1578 | -81.4046
3.1581 | -81.3962
3.1584 | -81.3862
3.1587 | -81.379
3.159 | -81.3721
3.1593 | -81.3649
3.1596 | -81.3595
3.1599 | -81.3532
3.1602 | -81.3487
3.1605 | -81.3469
3.1609 | -81.3426
3.1612 | -81.3367
3.1615 | -81.3304
3.1618 | -81.3242
3.1621 | -81.3193
3.1624 | -81.3011
3.1627 | -81.2945
3.163 | -81.287
3.1633 | -81.2814
3.1636 | -81.2756
3.1639 | -81.2699
3.1643 | -81.2646
3.1646 | -81.2592
3.1649 | -81.2523
3.1652 | -81.2444
3.1655 | -81.2343
3.1658 | -81.2284
3.1661 | -81.2228
3.1664 | -81.2152
3.1667 | -81.205
3.167 | -81.1946
3.1673 | -81.1872
3.1676 | -81.1797
3.168 | -81.1726
3.1683 | -81.1663
3.1686 | -81.1602
3.1689 | -81.152
3.1692 | -81.144
3.1695 | -81.1378
3.1698 | -81.1317
3.1701 | -81.1254
3.1704 | -81.118
3.1707 | -81.1117
3.171 | -81.1055
3.1714 | -81.0861
3.1717 | -81.0805
3.172 | -81.0707
3.1723 | -81.0637
3.1726 | -81.0559
3.1729 | -81.0483
3.1732 | -81.0425
3.1735 | -81.0353
3.1738 | -81.0282
3.1741 | -81.0202
3.1744 | -81.0135
3.1747 | -81.0077
3.1751 | -81.0008
3.1754 | -80.9929
3.1757 | -80.9857
3.176 | -80.9779
3.1763 | -80.9699
3.1766 | -80.962
3.1769 | -80.9524
3.1772 | -80.9429
3.1775 | -80.934
3.1778 | -80.9241
3.1781 | -80.9173
3.1785 | -80.9101
3.1788 | -80.9019
3.1791 | -80.8931
3.1794 | -80.8846
3.1797 | -80.8765
3.18 | -80.8676
3.1803 | -80.8557
3.1806 | -80.8465
3.1809 | -80.8369
3.1812 | -80.8279
3.1815 | -80.8202
3.1819 | -80.8122
3.1822 | -80.8034
3.1825 | -80.7943
3.1828 | -80.7854
3.1831 | -80.7782
3.1834 | -80.7709
3.1837 | -80.7622
3.184 | -80.7531
3.1843 | -80.7437
3.1846 | -80.7343
3.1849 | -80.7256
3.1852 | -80.7187
3.1856 | -80.7097
3.1859 | -80.6997
3.1862 | -80.6902
3.1865 | -80.681
3.1868 | -80.6717
3.1871 | -80.6635
3.1874 | -80.6548
3.1877 | -80.6457
3.188 | -80.6364
3.1883 | -80.626
3.1886 | -80.617
3.189 | -80.6096
3.1893 | -80.5956
3.1896 | -80.5858
3.1899 | -80.5756
3.1902 | -80.5657
3.1905 | -80.5546
3.1908 | -80.5437
3.1911 | -80.5341
3.1914 | -80.524
3.1917 | -80.514
3.192 | -80.504
3.1923 | -80.4941
3.1927 | -80.4834
3.193 | -80.4732
3.1933 | -80.4634
3.1936 | -80.4531
3.1939 | -80.4437
3.1942 | -80.4346
3.1945 | -80.4255
3.1948 | -80.4152
3.1951 | -80.4053
3.1954 | -80.3964
3.1957 | -80.3873
3.1961 | -80.3774
3.1964 | -80.3672
3.1967 | -80.3569
3.197 | -80.3469
3.1973 | -80.3378
3.1976 | -80.3278
3.1979 | -80.3182
3.1982 | -80.3049
3.1985 | -80.2956
3.1988 | -80.2851
3.1991 | -80.2749
3.1995 | -80.2648
3.1998 | -80.2549
3.2001 | -80.2452
3.2004 | -80.235
3.2007 | -80.2245
3.201 | -80.2142
3.2013 | -80.2041
3.2016 | -80.1941
3.2019 | -80.1831
3.2022 | -80.1715
3.2025 | -80.1619
3.2028 | -80.1534
3.2032 | -80.143
3.2035 | -80.133
3.2038 | -80.124
3.2041 | -80.1137
3.2044 | -80.1038
3.2047 | -80.0934
3.205 | -80.0827
3.2053 | -80.0706
3.2056 | -80.0593
3.2059 | -80.0486
3.2062 | -80.038
3.2066 | -80.0275
3.2069 | -80.0161
3.2072 | -80.0058
3.2075 | -79.9956
3.2078 | -79.9852
3.2081 | -79.9752
3.2084 | -79.965
3.2087 | -79.9552
3.209 | -79.9445
3.2093 | -79.934
3.2096 | -79.9238
3.21 | -79.9129
3.2103 | -79.9026
3.2106 | -79.8921
3.2109 | -79.8819
3.2112 | -79.8718
3.2115 | -79.8619
3.2118 | -79.8509
3.2121 | -79.8402
3.2124 | -79.8303
3.2127 | -79.8205
3.213 | -79.8105
3.2133 | -79.8005
3.2137 | -79.7907
3.214 | -79.7809
3.2143 | -79.7717
3.2146 | -79.7624
3.2149 | -79.7517
3.2152 | -79.7419
3.2155 | -79.732
3.2158 | -79.7235
3.2161 | -79.7139
3.2164 | -79.7041
3.2167 | -79.6948
3.2171 | -79.6861
3.2174 | -79.6764
3.2177 | -79.667
3.218 | -79.6575
3.2183 | -79.6483
3.2186 | -79.6389
3.2189 | -79.6293
3.2192 | -79.6196
3.2195 | -79.6105
3.2198 | -79.6014
3.2201 | -79.5941
3.2204 | -79.5848
3.2208 | -79.5751
3.2211 | -79.5664
3.2214 | -79.5577
3.2217 | -79.5483
3.222 | -79.5393
3.2223 | -79.5305
3.2226 | -79.522
3.2229 | -79.5133
3.2232 | -79.5052
3.2235 | -79.4976
3.2238 | -79.4894
3.2242 | -79.4803
3.2245 | -79.4717
3.2248 | -79.4657
3.2251 | -79.4563
3.2254 | -79.4483
3.2257 | -79.4398
3.226 | -79.4324
3.2263 | -79.4238
3.2266 | -79.4161
3.2269 | -79.408
3.2272 | -79.3999
3.2275 | -79.3916
3.2279 | -79.3833
3.2282 | -79.3751
3.2285 | -79.367
3.2288 | -79.3594
3.2291 | -79.3515
3.2294 | -79.3442
3.2297 | -79.3368
3.23 | -79.3292
3.2303 | -79.3212
3.2306 | -79.3129
3.2309 | -79.3045
3.2313 | -79.2973
3.2316 | -79.2898
3.2319 | -79.2826
3.2322 | -79.2759
3.2325 | -79.2683
3.2328 | -79.2614
3.2331 | -79.2547
3.2334 | -79.2477
3.2337 | -79.2447
3.234 | -79.2384
3.2343 | -79.2318
3.2347 | -79.2251
3.235 | -79.2196
3.2353 | -79.2137
3.2356 | -79.2072
3.2359 | -79.2011
3.2362 | -79.1948
3.2365 | -79.1884
3.2368 | -79.1822
3.2371 | -79.1761
3.2374 | -79.1698
3.2377 | -79.1643
3.238 | -79.1586
3.2384 | -79.1532
3.2387 | -79.1474
3.239 | -79.1417
3.2393 | -79.1362
3.2396 | -79.1306
3.2399 | -79.1251
3.2402 | -79.1203
3.2405 | -79.1155
3.2408 | -79.1101
3.2411 | -79.1045
3.2414 | -79.0994
3.2418 | -79.0947
3.2421 | -79.0899
3.2424 | -79.0854
3.2427 | -79.0844
3.243 | -79.0798
3.2433 | -79.075
3.2436 | -79.0699
3.2439 | -79.0657
3.2442 | -79.0617
3.2445 | -79.0574
3.2448 | -79.0532
3.2451 | -79.0495
3.2455 | -79.0445
3.2458 | -79.0385
3.2461 | -79.0334
3.2464 | -79.0295
3.2467 | -79.0258
3.247 | -79.0224
3.2473 | -79.0189
3.2476 | -79.0154
3.2479 | -79.0119
3.2482 | -79.0086
3.2485 | -79.0054
3.2489 | -79.0028
3.2492 | -79.0002
3.2495 | -78.9968
3.2498 | -78.995
3.2501 | -78.9941
3.2504 | -78.9909
3.2507 | -78.988
3.251 | -78.985
3.2513 | -78.983
3.2516 | -78.9866
3.2519 | -78.9848
3.2523 | -78.9825
3.2526 | -78.9801
3.2529 | -78.9775
3.2532 | -78.9752
3.2535 | -78.9731
3.2538 | -78.9711
3.2541 | -78.9693
3.2544 | -78.9684
3.2547 | -78.9683
3.255 | -78.9667
3.2553 | -78.9653
3.2556 | -78.9643
3.256 | -78.9635
3.2563 | -78.9621
3.2566 | -78.9613
3.2569 | -78.9606
3.2572 | -78.9599
3.2575 | -78.9588
3.2578 | -78.9586
3.2581 | -78.9594
3.2584 | -78.959
3.2587 | -78.9586
3.259 | -78.9585
3.2594 | -78.9589
3.2597 | -78.9592
3.26 | -78.9592
3.2603 | -78.9654
3.2606 | -78.966
3.2609 | -78.9668
3.2612 | -78.9676
3.2615 | -78.9698
3.2618 | -78.972
3.2621 | -78.9725
3.2624 | -78.9738
3.2627 | -78.9748
3.2631 | -78.9759
3.2634 | -78.9767
3.2637 | -78.9778
3.264 | -78.9791
3.2643 | -78.9805
3.2646 | -78.982
3.2649 | -78.9833
3.2652 | -78.9851
3.2655 | -78.9871
3.2658 | -78.9891
3.2661 | -78.9912
3.2665 | -78.9936
3.2668 | -78.9956
3.2671 | -78.9974
3.2674 | -78.9997
3.2677 | -79.0027
3.268 | -79.0051
3.2683 | -79.0073
3.2686 | -79.0096
3.2689 | -79.0121
3.2692 | -79.0205
3.2695 | -79.0229
3.2699 | -79.0255
3.2702 | -79.0284
3.2705 | -79.0314
3.2708 | -79.0346
3.2711 | -79.0377
3.2714 | -79.0413
3.2717 | -79.0447
3.272 | -79.0481
3.2723 | -79.0515
3.2726 | -79.0555
3.2729 | -79.0594
3.2732 | -79.063
3.2736 | -79.0672
3.2739 | -79.0714
3.2742 | -79.0754
3.2745 | -79.0794
3.2748 | -79.0837
3.2751 | -79.0888
3.2754 | -79.0937
3.2757 | -79.0983
3.276 | -79.1027
3.2763 | -79.1074
3.2766 | -79.1123
3.277 | -79.1179
3.2773 | -79.1234
3.2776 | -79.1288
3.2779 | -79.1341
3.2782 | -79.1449
3.2785 | -79.1502
3.2788 | -79.1558
3.2791 | -79.1617
3.2794 | -79.1682
3.2797 | -79.1743
3.28 | -79.1811
3.2803 | -79.189
3.2807 | -79.1964
3.281 | -79.2025
3.2813 | -79.208
3.2816 | -79.2139
3.2819 | -79.2195
3.2822 | -79.2252
3.2825 | -79.2318
3.2828 | -79.2384
3.2831 | -79.2444
3.2834 | -79.2507
3.2837 | -79.2577
3.2841 | -79.2653
3.2844 | -79.2709
3.2847 | -79.2766
3.285 | -79.2832
3.2853 | -79.2899
3.2856 | -79.297
3.2859 | -79.3043
3.2862 | -79.3114
3.2865 | -79.3187
3.2868 | -79.3266
3.2871 | -79.3386
3.2875 | -79.3469
3.2878 | -79.3556
3.2881 | -79.3639
3.2884 | -79.3721
3.2887 | -79.3801
3.289 | -79.3882
3.2893 | -79.3966
3.2896 | -79.4052
3.2899 | -79.4136
3.2902 | -79.4224
3.2905 | -79.432
3.2908 | -79.4405
3.2912 | -79.4493
3.2915 | -79.4577
3.2918 | -79.4658
3.2921 | -79.4739
3.2924 | -79.4824
3.2927 | -79.4927
3.293 | -79.5033
3.2933 | -79.5125
3.2936 | -79.5219
3.2939 | -79.531
3.2942 | -79.5406
3.2946 | -79.5499
3.2949 | -79.5598
3.2952 | -79.5695
3.2955 | -79.5795
3.2958 | -79.5892
3.2961 | -79.6008
3.2964 | -79.6094
3.2967 | -79.619
3.297 | -79.6295
3.2973 | -79.6399
3.2976 | -79.6496
3.2979 | -79.6597
3.2983 | -79.6693
3.2986 | -79.6789
3.2989 | -79.6887
3.2992 | -79.6988
3.2995 | -79.7093
3.2998 | -79.7195
3.3001 | -79.7297
3.3004 | -79.7401
3.3007 | -79.7502
3.301 | -79.7605
3.3013 | -79.7711
3.3017 | -79.7815
3.302 | -79.7915
3.3023 | -79.8022
3.3026 | -79.8128
3.3029 | -79.8233
3.3032 | -79.8344
3.3035 | -79.846
3.3038 | -79.8572
3.3041 | -79.8676
3.3044 | -79.8787
3.3047 | -79.89
3.3051 | -79.903
3.3054 | -79.9142
3.3057 | -79.9251
3.306 | -79.9362
3.3063 | -79.9482
3.3066 | -79.9596
3.3069 | -79.9714
3.3072 | -79.9829
3.3075 | -79.9948
3.3078 | -80.0065
3.3081 | -80.0184
3.3084 | -80.0296
3.3088 | -80.0413
3.3091 | -80.0532
3.3094 | -80.0651
3.3097 | -80.0772
3.31 | -80.0892
3.3103 | -80.1029
3.3106 | -80.117
3.3109 | -80.1298
3.3112 | -80.1421
3.3115 | -80.1542
3.3118 | -80.1672
3.3122 | -80.1803
3.3125 | -80.1929
3.3128 | -80.2052
3.3131 | -80.2178
3.3134 | -80.2301
3.3137 | -80.2425
3.314 | -80.2549
3.3143 | -80.2673
3.3146 | -80.2803
3.3149 | -80.2931
3.3152 | -80.3058
3.3155 | -80.3184
3.3159 | -80.3304
3.3162 | -80.343
3.3165 | -80.3559
3.3168 | -80.3682
3.3171 | -80.3808
3.3174 | -80.3936
3.3177 | -80.4071
3.318 | -80.4199
3.3183 | -80.4328
3.3186 | -80.4455
3.3189 | -80.4584
3.3193 | -80.4715
3.3196 | -80.4847
3.3199 | -80.4981
3.3202 | -80.5113
3.3205 | -80.5243
3.3208 | -80.5374
3.3211 | -80.5506
3.3214 | -80.5642
3.3217 | -80.5783
3.322 | -80.5921
3.3223 | -80.6073
3.3226 | -80.6195
3.323 | -80.6326
3.3233 | -80.6457
3.3236 | -80.6599
3.3239 | -80.6729
3.3242 | -80.6859
3.3245 | -80.6986
3.3248 | -80.7119
3.3251 | -80.7252
3.3254 | -80.7387
3.3257 | -80.7521
3.326 | -80.7655
3.3264 | -80.7789
3.3267 | -80.7929
3.327 | -80.8072
3.3273 | -80.8189
3.3276 | -80.8294
3.3279 | -80.8414
3.3282 | -80.8558
3.3285 | -80.8699
3.3288 | -80.8827
3.3291 | -80.8954
3.3294 | -80.909
3.3298 | -80.9225
3.3301 | -80.9356
3.3304 | -80.9488
3.3307 | -80.962
3.331 | -80.9751
3.3313 | -80.9881
3.3316 | -80.9989
3.3319 | -81.0122
3.3322 | -81.0263
3.3325 | -81.0394
3.3328 | -81.0521
3.3331 | -81.0655
3.3335 | -81.0789
3.3338 | -81.0905
3.3341 | -81.1033
3.3344 | -81.1169
3.3347 | -81.1297
3.335 | -81.1416
3.3353 | -81.1538
3.3356 | -81.1681
3.3359 | -81.1812
3.3362 | -81.1945
3.3365 | -81.2074
3.3369 | -81.2201
3.3372 | -81.2338
3.3375 | -81.2458
3.3378 | -81.2581
3.3381 | -81.2707
3.3384 | -81.2829
3.3387 | -81.2946
3.339 | -81.3069
3.3393 | -81.3206
3.3396 | -81.3342
3.3399 | -81.3474
3.3403 | -81.3614
3.3406 | -81.3701
3.3409 | -81.3843
3.3412 | -81.3971
3.3415 | -81.4094
3.3418 | -81.4223
3.3421 | -81.4355
3.3424 | -81.4475
3.3427 | -81.4591
3.343 | -81.4715
3.3433 | -81.4838
3.3436 | -81.4952
3.344 | -81.5072
3.3443 | -81.5196
3.3446 | -81.5313
3.3449 | -81.5442
3.3452 | -81.5563
3.3455 | -81.5705
3.3458 | -81.5843
3.3461 | -81.597
3.3464 | -81.6089
3.3467 | -81.6206
3.347 | -81.633
3.3474 | -81.6463
3.3477 | -81.6603
3.348 | -81.6722
3.3483 | -81.6848
3.3486 | -81.6967
3.3489 | -81.7083
3.3492 | -81.72
3.3495 | -81.7268
3.3498 | -81.739
3.3501 | -81.7506
3.3504 | -81.7622
3.3507 | -81.7735
3.3511 | -81.7841
3.3514 | -81.796
3.3517 | -81.8081
3.352 | -81.82
3.3523 | -81.8324
3.3526 | -81.8437
3.3529 | -81.8548
3.3532 | -81.8652
3.3535 | -81.8779
3.3538 | -81.8938
3.3541 | -81.9046
3.3545 | -81.9154
3.3548 | -81.9262
3.3551 | -81.939
3.3554 | -81.9498
3.3557 | -81.9608
3.356 | -81.9747
3.3563 | -81.9849
3.3566 | -81.9974
3.3569 | -82.0095
3.3572 | -82.0199
3.3575 | -82.0293
3.3578 | -82.0392
3.3582 | -82.049
3.3585 | -82.0542
3.3588 | -82.064
3.3591 | -82.0743
3.3594 | -82.0838
3.3597 | -82.0942
3.36 | -82.1046
3.3603 | -82.1135
3.3606 | -82.1227
3.3609 | -82.1329
3.3612 | -82.1435
3.3616 | -82.1514
3.3619 | -82.1603
3.3622 | -82.1696
3.3625 | -82.1798
3.3628 | -82.1898
3.3631 | -82.1992
3.3634 | -82.2109
3.3637 | -82.2225
3.364 | -82.2325
3.3643 | -82.2424
3.3646 | -82.252
3.365 | -82.262
3.3653 | -82.2714
3.3656 | -82.2827
3.3659 | -82.2924
3.3662 | -82.3
3.3665 | -82.3096
3.3668 | -82.3216
3.3671 | -82.3375
3.3674 | -82.3422
3.3677 | -82.3552
3.368 | -82.3625
3.3683 | -82.3722
3.3687 | -82.3826
3.369 | -82.3939
3.3693 | -82.4065
3.3696 | -82.4183
3.3699 | -82.4277
3.3702 | -82.4337
3.3705 | -82.4401
3.3708 | -82.4474
3.3711 | -82.4561
3.3714 | -82.4709
3.3717 | -82.4858
3.3721 | -82.4955
3.3724 | -82.503
3.3727 | -82.5101
3.373 | -82.5182
3.3733 | -82.5305
3.3736 | -82.5407
3.3739 | -82.55
3.3742 | -82.5596
3.3745 | -82.5689
3.3748 | -82.5759
3.3751 | -82.583
3.3754 | -82.5911
3.3758 | -82.5975
3.3761 | -82.6071
3.3764 | -82.6101
3.3767 | -82.6202
3.377 | -82.6259
3.3773 | -82.6367
3.3776 | -82.6465
3.3779 | -82.6547
3.3782 | -82.6628
3.3785 | -82.6721
3.3788 | -82.6827
3.3792 | -82.698
3.3795 | -82.7097
3.3798 | -82.7153
3.3801 | -82.7248
3.3804 | -82.7329
3.3807 | -82.738
3.381 | -82.7424
3.3813 | -82.7485
3.3816 | -82.7537
3.3819 | -82.7578
3.3822 | -82.767
3.3826 | -82.7827
3.3829 | -82.7947
3.3832 | -82.801
3.3835 | -82.8113
3.3838 | -82.819
3.3841 | -82.8234
3.3844 | -82.8389
3.3847 | -82.8566
3.385 | -82.8569
3.3853 | -82.8647
3.3856 | -82.8742
3.3859 | -82.8856
3.3863 | -82.8922
3.3866 | -82.8996
3.3869 | -82.9068
3.3872 | -82.9106
3.3875 | -82.9163
3.3878 | -82.9218
3.3881 | -82.9254
3.3884 | -82.9291
3.3887 | -82.9349
3.389 | -82.9397
3.3893 | -82.9444
3.3897 | -82.9474
3.39 | -82.9576
3.3903 | -82.9662
3.3906 | -82.9691
3.3909 | -82.9734
3.3912 | -82.9757
3.3915 | -82.9792
3.3918 | -82.9832
3.3921 | -82.985
3.3924 | -82.9872
3.3927 | -82.9913
3.393 | -82.9965
3.3934 | -83.0055
3.3937 | -83.0085
3.394 | -83.0093
3.3943 | -83.0106
3.3946 | -83.016
3.3949 | -83.0232
3.3952 | -83.0319
3.3955 | -83.0342
3.3958 | -83.0362
3.3961 | -83.0397
3.3964 | -83.0414
3.3968 | -83.0419
3.3971 | -83.0447
3.3974 | -83.048
3.3977 | -83.0501
3.398 | -83.0527
3.3983 | -83.0552
3.3986 | -83.0574
3.3989 | -83.0589
3.3992 | -83.0616
3.3995 | -83.0697
3.3998 | -83.0766
3.4002 | -83.0773
3.4005 | -83.0846
3.4008 | -83.0878
3.4011 | -83.0888
3.4014 | -83.0913
3.4017 | -83.0926
3.402 | -83.0942
3.4023 | -83.0946
3.4026 | -83.0965
3.4029 | -83.098
3.4032 | -83.0998
3.4035 | -83.1022
3.4039 | -83.1044
3.4042 | -83.1055
3.4045 | -83.1056
3.4048 | -83.1044
3.4051 | -83.1044
3.4054 | -83.1053
3.4057 | -83.1085
3.406 | -83.1098
3.4063 | -83.1109
3.4066 | -83.1142
3.4069 | -83.1143
3.4073 | -83.1154
3.4076 | -83.1167
3.4079 | -83.119
3.4082 | -83.1185
3.4085 | -83.1194
3.4088 | -83.1212
3.4091 | -83.1217
3.4094 | -83.1215
3.4097 | -83.1237
3.41 | -83.1267
3.4103 | -83.1282
3.4106 | -83.1293
3.411 | -83.1308
3.4113 | -83.1313
3.4116 | -83.1333
3.4119 | -83.1349
3.4122 | -83.1359
3.4125 | -83.1364
3.4128 | -83.1376
3.4131 | -83.1382
3.4134 | -83.1391
3.4137 | -83.1399
3.414 | -83.1402
3.4144 | -83.1416
3.4147 | -83.1435
3.415 | -83.1444
3.4153 | -83.1461
3.4156 | -83.1465
3.4159 | -83.1463
3.4162 | -83.1469
3.4165 | -83.1476
3.4168 | -83.1485
3.4171 | -83.1492
3.4174 | -83.1501
3.4177 | -83.1519
3.4181 | -83.1537
3.4184 | -83.1542
3.4187 | -83.1549
3.419 | -83.1573
3.4193 | -83.1583
3.4196 | -83.1596
3.4199 | -83.1622
3.4202 | -83.1634
3.4205 | -83.1633
3.4208 | -83.1641
3.4211 | -83.1653
3.4215 | -83.1669
3.4218 | -83.1684
3.4221 | -83.1684
3.4224 | -83.1685
3.4227 | -83.1704
3.423 | -83.1726
3.4233 | -83.1718
3.4236 | -83.1716
3.4239 | -83.1722
3.4242 | -83.1741
3.4245 | -83.1748
3.4249 | -83.1758
3.4252 | -83.1768
3.4255 | -83.1778
3.4258 | -83.1785
3.4261 | -83.1788
3.4264 | -83.1789
3.4267 | -83.1804
3.427 | -83.1799
3.4273 | -83.1806
3.4276 | -83.1816
3.4279 | -83.1823
3.4282 | -83.1834
3.4286 | -83.185
3.4289 | -83.1849
3.4292 | -83.1852
3.4295 | -83.1864
3.4298 | -83.186
3.4301 | -83.1855
3.4304 | -83.1851
3.4307 | -83.1869
3.431 | -83.1872
3.4313 | -83.1881
3.4316 | -83.1884
3.432 | -83.1874
3.4323 | -83.1867
3.4326 | -83.188
3.4329 | -83.1905
3.4332 | -83.1907
3.4335 | -83.1916
3.4338 | -83.1923
3.4341 | -83.1919
3.4344 | -83.1924
3.4347 | -83.1932
3.435 | -83.193
3.4354 | -83.1924
3.4357 | -83.1934
3.436 | -83.198
3.4363 | -83.2019
3.4366 | -83.2029
3.4369 | -83.2038
3.4372 | -83.2049
3.4375 | -83.2062
3.4378 | -83.2076
3.4381 | -83.2087
3.4384 | -83.2096
3.4387 | -83.2094
3.4391 | -83.2097
3.4394 | -83.209
3.4397 | -83.2087
3.44 | -83.2092
3.4403 | -83.2107
3.4406 | -83.2118
3.4409 | -83.211
3.4412 | -83.2132
3.4415 | -83.2145
3.4418 | -83.2153
3.4421 | -83.2161
3.4425 | -83.2159
3.4428 | -83.2168
3.4431 | -83.219
3.4434 | -83.2205
3.4437 | -83.2212
3.444 | -83.2217
3.4443 | -83.2227
3.4446 | -83.2234
3.4449 | -83.2243
3.4452 | -83.2251
3.4455 | -83.2276
3.4458 | -83.2304
3.4462 | -83.2308
3.4465 | -83.2298
3.4468 | -83.2285
3.4471 | -83.2266
3.4474 | -83.2226
3.4477 | -83.223
3.448 | -83.2234
3.4483 | -83.2242
3.4486 | -83.2244
3.4489 | -83.2252
3.4492 | -83.2258
3.4496 | -83.2251
3.4499 | -83.2264
3.4502 | -83.2268
3.4505 | -83.2265
3.4508 | -83.2265
3.4511 | -83.2275
3.4514 | -83.2278
3.4517 | -83.2282
3.452 | -83.2285
3.4523 | -83.2296
3.4526 | -83.232
3.4529 | -83.2322
3.4533 | -83.2325
3.4536 | -83.2334
3.4539 | -83.234
3.4542 | -83.2347
3.4545 | -83.2352
3.4548 | -83.2362
3.4551 | -83.2373
3.4554 | -83.2372
3.4557 | -83.2375
3.456 | -83.2391
3.4563 | -83.2404
3.4567 | -83.239
3.457 | -83.2379
3.4573 | -83.239
3.4576 | -83.2419
3.4579 | -83.2432
3.4582 | -83.245
3.4585 | -83.2455
3.4588 | -83.247
3.4591 | -83.2481
3.4594 | -83.2484
3.4597 | -83.2496
3.4601 | -83.2505
3.4604 | -83.2518
3.4607 | -83.2526
3.461 | -83.2524
3.4613 | -83.2493
3.4616 | -83.2479
3.4619 | -83.2487
3.4622 | -83.2487
3.4625 | -83.2492
3.4628 | -83.251
3.4631 | -83.2542
3.4634 | -83.257
3.4638 | -83.2581
3.4641 | -83.2594
3.4644 | -83.2603
3.4647 | -83.2611
3.465 | -83.263
3.4653 | -83.2606
3.4656 | -83.2606
3.4659 | -83.2612
3.4662 | -83.2615
3.4665 | -83.262
3.4668 | -83.2623
3.4672 | -83.2628
3.4675 | -83.2629
3.4678 | -83.2635
3.4681 | -83.2647
3.4684 | -83.2655
3.4687 | -83.2661
3.469 | -83.2667
3.4693 | -83.2676
3.4696 | -83.2663
3.4699 | -83.2653
3.4702 | -83.2649
3.4705 | -83.2653
3.4709 | -83.2658
3.4712 | -83.2664
3.4715 | -83.2668
3.4718 | -83.2657
3.4721 | -83.2638
3.4724 | -83.2634
3.4727 | -83.2641
3.473 | -83.2635
3.4733 | -83.2637
3.4736 | -83.2639
3.4739 | -83.2648
3.4743 | -83.2651
3.4746 | -83.2667
3.4749 | -83.2677
3.4752 | -83.2677
3.4755 | -83.266
3.4758 | -83.2666
3.4761 | -83.2676
3.4764 | -83.2693
3.4767 | -83.2692
3.477 | -83.2704
3.4773 | -83.2724
3.4777 | -83.2738
3.478 | -83.2747
3.4783 | -83.2755
3.4786 | -83.2757
3.4789 | -83.2764
3.4792 | -83.2791
3.4795 | -83.2848
3.4798 | -83.288
3.4801 | -83.2891
3.4804 | -83.2894
3.4807 | -83.2897
3.481 | -83.2899
3.4814 | -83.2903
3.4817 | -83.2901
3.482 | -83.2899
3.4823 | -83.2907
3.4826 | -83.2916
3.4829 | -83.292
3.4832 | -83.2932
3.4835 | -83.2953
3.4838 | -83.2963
3.4841 | -83.2958
3.4844 | -83.2974
3.4848 | -83.2986
3.4851 | -83.2991
3.4854 | -83.2996
3.4857 | -83.2993
3.486 | -83.2994
3.4863 | -83.299
3.4866 | -83.2997
3.4869 | -83.3007
3.4872 | -83.3028
3.4875 | -83.3048
3.4878 | -83.3058
3.4881 | -83.306
3.4885 | -83.3057
3.4888 | -83.3052
3.4891 | -83.3038
3.4894 | -83.3029
3.4897 | -83.3027
3.49 | -83.3025
3.4903 | -83.3022
3.4906 | -83.3028
3.4909 | -83.3033
3.4912 | -83.3035
3.4915 | -83.3036
3.4919 | -83.3046
3.4922 | -83.3051
3.4925 | -83.3059
3.4928 | -83.3058
3.4931 | -83.3044
3.4934 | -83.3029
3.4937 | -83.3025
3.494 | -83.3029
3.4943 | -83.3031
3.4946 | -83.3033
3.4949 | -83.3034
3.4953 | -83.3019
3.4956 | -83.3014
3.4959 | -83.3016
3.4962 | -83.3017
3.4965 | -83.302
3.4968 | -83.3019
3.4971 | -83.3013
3.4974 | -83.3014
3.4977 | -83.302
3.498 | -83.3011
3.4983 | -83.2994
3.4986 | -83.2995
3.499 | -83.2997
3.4993 | -83.2995
3.4996 | -83.2989
3.4999 | -83.2991
3.5002 | -83.3
3.5005 | -83.3016
3.5008 | -83.3024
3.5011 | -83.3031
3.5014 | -83.3039
3.5017 | -83.3021
3.502 | -83.2976
3.5024 | -83.2973
3.5027 | -83.2974
3.503 | -83.2977
3.5033 | -83.2976
3.5036 | -83.2971
3.5039 | -83.2973
3.5042 | -83.2971
3.5045 | -83.2979
3.5048 | -83.2979
3.5051 | -83.299
3.5054 | -83.299
3.5057 | -83.2977
3.5061 | -83.2978
3.5064 | -83.2979
3.5067 | -83.298
3.507 | -83.2971
3.5073 | -83.2969
3.5076 | -83.2965
3.5079 | -83.297
3.5082 | -83.297
3.5085 | -83.2972
3.5088 | -83.2976
3.5091 | -83.2972
3.5095 | -83.2963
3.5098 | -83.2967
3.5101 | -83.2972
3.5104 | -83.2977
3.5107 | -83.2975
3.511 | -83.298
3.5113 | -83.2955
3.5116 | -83.2933
3.5119 | -83.2932
3.5122 | -83.2926
3.5125 | -83.2925
3.5129 | -83.2925
3.5132 | -83.2923
3.5135 | -83.2919
3.5138 | -83.2915
3.5141 | -83.2911
3.5144 | -83.2916
3.5147 | -83.2909
3.515 | -83.2899
3.5153 | -83.2902
3.5156 | -83.2904
3.5159 | -83.2903
3.5162 | -83.2915
3.5166 | -83.2919
3.5169 | -83.2921
3.5172 | -83.2927
3.5175 | -83.293
3.5178 | -83.2925
3.5181 | -83.2935
3.5184 | -83.2946
3.5187 | -83.2952
3.519 | -83.2948
3.5193 | -83.294
3.5196 | -83.2948
3.52 | -83.2942
3.5203 | -83.2941
3.5206 | -83.2939
3.5209 | -83.2968
3.5212 | -83.2992
3.5215 | -83.3
3.5218 | -83.3008
3.5221 | -83.3013
3.5224 | -83.3013
3.5227 | -83.3018
3.523 | -83.3018
3.5233 | -83.3021
3.5237 | -83.3023
3.524 | -83.303
3.5243 | -83.3038
3.5246 | -83.3031
3.5249 | -83.3027
3.5252 | -83.303
3.5255 | -83.3032
3.5258 | -83.3031
3.5261 | -83.302
3.5264 | -83.3003
3.5267 | -83.2993
3.5271 | -83.2987
3.5274 | -83.2983
3.5277 | -83.2979
3.528 | -83.2971
3.5283 | -83.2974
3.5286 | -83.2978
3.5289 | -83.2984
3.5292 | -83.299
3.5295 | -83.2994
3.5298 | -83.2999
3.5301 | -83.3003
3.5305 | -83.2997
3.5308 | -83.3
3.5311 | -83.3013
3.5314 | -83.3024
3.5317 | -83.3021
3.532 | -83.3015
3.5323 | -83.3014
3.5326 | -83.2978
3.5329 | -83.293
3.5332 | -83.293
3.5335 | -83.2954
3.5338 | -83.2958
3.5342 | -83.2961
3.5345 | -83.2961
3.5348 | -83.2946
3.5351 | -83.2927
3.5354 | -83.2924
3.5357 | -83.2927
3.536 | -83.2934
3.5363 | -83.2941
3.5366 | -83.2941
3.5369 | -83.2931
3.5372 | -83.2923
3.5376 | -83.2926
3.5379 | -83.2926
3.5382 | -83.2916
3.5385 | -83.2915
3.5388 | -83.2911
3.5391 | -83.2907
3.5394 | -83.2913
3.5397 | -83.2922
3.54 | -83.2915
3.5403 | -83.2915
3.5406 | -83.2918
3.5409 | -83.2914
3.5413 | -83.291
3.5416 | -83.29
3.5419 | -83.2901
3.5422 | -83.2899
3.5425 | -83.2897
3.5428 | -83.2899
3.5431 | -83.2896
3.5434 | -83.2894
3.5437 | -83.2898
3.544 | -83.2899
3.5443 | -83.2897
3.5447 | -83.289
3.545 | -83.2884
3.5453 | -83.2876
3.5456 | -83.2884
3.5459 | -83.2873
3.5462 | -83.2872
3.5465 | -83.2864
3.5468 | -83.2868
3.5471 | -83.2872
3.5474 | -83.2866
3.5477 | -83.2862
3.548 | -83.2862
3.5484 | -83.2863
3.5487 | -83.2862
3.549 | -83.2835
3.5493 | -83.2791
3.5496 | -83.2765
3.5499 | -83.2759
3.5502 | -83.275
3.5505 | -83.2746
3.5508 | -83.2745
3.5511 | -83.2739
3.5514 | -83.2733
3.5518 | -83.2728
3.5521 | -83.2721
3.5524 | -83.2719
3.5527 | -83.2726
3.553 | -83.2727
3.5533 | -83.2723
3.5536 | -83.2725
3.5539 | -83.2723
3.5542 | -83.2726
3.5545 | -83.2734
3.5548 | -83.2754
3.5552 | -83.2756
3.5555 | -83.2749
3.5558 | -83.2753
3.5561 | -83.2756
3.5564 | -83.2753
3.5567 | -83.2755
3.557 | -83.2753
3.5573 | -83.2754
3.5576 | -83.275
3.5579 | -83.274
3.5582 | -83.2734
3.5585 | -83.2733
3.5589 | -83.2736
3.5592 | -83.2738
3.5595 | -83.2734
3.5598 | -83.2726
3.5601 | -83.2721
3.5604 | -83.2715
3.5607 | -83.2715
3.561 | -83.2713
3.5613 | -83.2691
3.5616 | -83.2673
3.5619 | -83.2662
3.5623 | -83.2676
3.5626 | -83.2681
3.5629 | -83.2679
3.5632 | -83.2686
3.5635 | -83.2684
3.5638 | -83.2684
3.5641 | -83.2695
3.5644 | -83.27
3.5647 | -83.2699
3.565 | -83.2704
3.5653 | -83.2709
3.5657 | -83.2709
3.566 | -83.2709
3.5663 | -83.2711
3.5666 | -83.2714
3.5669 | -83.2737
3.5672 | -83.2777
3.5675 | -83.2786
3.5678 | -83.2784
3.5681 | -83.2784
3.5684 | -83.2796
3.5687 | -83.279
3.569 | -83.2783
3.5694 | -83.2778
3.5697 | -83.2777
3.57 | -83.2769
3.5703 | -83.2762
3.5706 | -83.2761
3.5709 | -83.2749
3.5712 | -83.2727
3.5715 | -83.2726
3.5718 | -83.2725
3.5721 | -83.2739
3.5724 | -83.2742
3.5728 | -83.275
3.5731 | -83.2756
3.5734 | -83.2754
3.5737 | -83.2759
3.574 | -83.2783
3.5743 | -83.2805
3.5746 | -83.2806
3.5749 | -83.2798
3.5752 | -83.2799
3.5755 | -83.2802
3.5758 | -83.2783
3.5761 | -83.2759
3.5765 | -83.2743
3.5768 | -83.2744
3.5771 | -83.2748
3.5774 | -83.2745
3.5777 | -83.2768
3.578 | -83.2778
3.5783 | -83.2777
3.5786 | -83.2781
3.5789 | -83.279
3.5792 | -83.2794
3.5795 | -83.2794
3.5799 | -83.2793
3.5802 | -83.279
3.5805 | -83.2791
3.5808 | -83.2793
3.5811 | -83.2806
3.5814 | -83.2814
3.5817 | -83.2816
3.582 | -83.2822
3.5823 | -83.2824
3.5826 | -83.2823
3.5829 | -83.2825
3.5833 | -83.283
3.5836 | -83.2834
3.5839 | -83.2835
3.5842 | -83.2833
3.5845 | -83.2843
3.5848 | -83.2856
3.5851 | -83.2863
3.5854 | -83.2873
3.5857 | -83.2875
3.586 | -83.2883
3.5863 | -83.2885
3.5866 | -83.2885
3.587 | -83.2889
3.5873 | -83.2895
3.5876 | -83.2895
3.5879 | -83.29
3.5882 | -83.2903
3.5885 | -83.2902
3.5888 | -83.2904
3.5891 | -83.2901
3.5894 | -83.2902
3.5897 | -83.2903
3.59 | -83.2911
3.5904 | -83.29
3.5907 | -83.2897
3.591 | -83.2909
3.5913 | -83.2912
3.5916 | -83.2911
3.5919 | -83.2912
3.5922 | -83.2917
3.5925 | -83.2927
3.5928 | -83.2932
3.5931 | -83.2932
3.5934 | -83.2924
3.5937 | -83.2915
3.5941 | -83.2896
3.5944 | -83.2892
3.5947 | -83.2908
3.595 | -83.2916
3.5953 | -83.2918
3.5956 | -83.2938
3.5959 | -83.294
3.5962 | -83.2945
3.5965 | -83.2947
3.5968 | -83.2953
3.5971 | -83.2966
3.5975 | -83.2969
3.5978 | -83.2972
3.5981 | -83.2965
3.5984 | -83.2974
3.5987 | -83.2995
3.599 | -83.3011
3.5993 | -83.3023
3.5996 | -83.3029
3.5999 | -83.3037
3.6002 | -83.3049
3.6005 | -83.3056
3.6008 | -83.3063
3.6012 | -83.3075
3.6015 | -83.309
3.6018 | -83.3095
3.6021 | -83.3093
3.6024 | -83.3102
3.6027 | -83.3109
3.603 | -83.3106
3.6033 | -83.3114
3.6036 | -83.3139
3.6039 | -83.3158
3.6042 | -83.3142
3.6046 | -83.3133
3.6049 | -83.3143
3.6052 | -83.3147
3.6055 | -83.3151
3.6058 | -83.3155
3.6061 | -83.3157
3.6064 | -83.3158
3.6067 | -83.3171
3.607 | -83.3174
3.6073 | -83.3179
3.6076 | -83.3182
3.608 | -83.3196
3.6083 | -83.3218
3.6086 | -83.3233
3.6089 | -83.3246
3.6092 | -83.3244
3.6095 | -83.3239
3.6098 | -83.3233
3.6101 | -83.3228
3.6104 | -83.3233
3.6107 | -83.3247
3.611 | -83.3261
3.6113 | -83.3267
3.6117 | -83.3268
3.612 | -83.3259
3.6123 | -83.3255
3.6126 | -83.325
3.6129 | -83.3274
3.6132 | -83.3311
3.6135 | -83.3318
3.6138 | -83.332
3.6141 | -83.3324
3.6144 | -83.3325
3.6147 | -83.3324
3.6151 | -83.3332
3.6154 | -83.3339
3.6157 | -83.3346
3.616 | -83.3344
3.6163 | -83.3338
3.6166 | -83.3335
3.6169 | -83.332
3.6172 | -83.3324
3.6175 | -83.3336
3.6178 | -83.336
3.6181 | -83.3376
3.6184 | -83.3384
3.6188 | -83.3393
3.6191 | -83.3408
3.6194 | -83.3418
3.6197 | -83.3422
3.62 | -83.3425
3.6203 | -83.3433
3.6206 | -83.3437
3.6209 | -83.344
3.6212 | -83.345
3.6215 | -83.3452
3.6218 | -83.3448
3.6222 | -83.3456
3.6225 | -83.3468
3.6228 | -83.3474
3.6231 | -83.3483
3.6234 | -83.3493
3.6237 | -83.3499
3.624 | -83.3505
3.6243 | -83.3521
3.6246 | -83.3536
3.6249 | -83.3542
3.6252 | -83.3558
3.6256 | -83.3569
3.6259 | -83.357
3.6262 | -83.3561
3.6265 | -83.3563
3.6268 | -83.3569
3.6271 | -83.358
3.6274 | -83.3585
3.6277 | -83.3584
3.628 | -83.3591
3.6283 | -83.3596
3.6286 | -83.3598
3.6289 | -83.3624
3.6293 | -83.3627
3.6296 | -83.3632
3.6299 | -83.3641
3.6302 | -83.366
3.6305 | -83.3669
3.6308 | -83.3674
3.6311 | -83.3688
3.6314 | -83.3701
3.6317 | -83.3711
3.632 | -83.3722
3.6323 | -83.3731
3.6327 | -83.3738
3.633 | -83.3753
3.6333 | -83.3758
3.6336 | -83.3765
3.6339 | -83.3777
3.6342 | -83.3793
3.6345 | -83.3813
3.6348 | -83.3846
3.6351 | -83.3868
3.6354 | -83.3861
3.6357 | -83.3855
3.636 | -83.3863
3.6364 | -83.3875
3.6367 | -83.3888
3.637 | -83.3896
3.6373 | -83.3912
3.6376 | -83.3926
3.6379 | -83.3938
3.6382 | -83.3943
3.6385 | -83.3953
3.6388 | -83.3971
3.6391 | -83.3982
3.6394 | -83.3988
3.6398 | -83.3996
3.6401 | -83.4018
3.6404 | -83.4026
3.6407 | -83.4032
3.641 | -83.4022
3.6413 | -83.404
3.6416 | -83.4061
3.6419 | -83.4073
3.6422 | -83.4078
3.6425 | -83.4077
3.6428 | -83.4076
3.6432 | -83.4077
3.6435 | -83.4095
3.6438 | -83.4105
3.6441 | -83.4119
3.6444 | -83.4136
3.6447 | -83.4149
3.645 | -83.4152
3.6453 | -83.4153
3.6456 | -83.4164
3.6459 | -83.4176
3.6462 | -83.4177
3.6465 | -83.4172
3.6469 | -83.4181
3.6472 | -83.4192
3.6475 | -83.4202
3.6478 | -83.4209
3.6481 | -83.4221
3.6484 | -83.4231
3.6487 | -83.4243
3.649 | -83.4255
3.6493 | -83.4262
3.6496 | -83.4273
3.6499 | -83.4307
3.6503 | -83.4332
3.6506 | -83.4344
3.6509 | -83.4354
3.6512 | -83.4378
3.6515 | -83.4394
3.6518 | -83.4405
3.6521 | -83.4421
3.6524 | -83.4437
3.6527 | -83.4447
3.653 | -83.4453
3.6533 | -83.4449
3.6536 | -83.4468
3.654 | -83.4475
3.6543 | -83.4483
3.6546 | -83.4501
3.6549 | -83.4523
3.6552 | -83.4533
3.6555 | -83.4545
3.6558 | -83.455
3.6561 | -83.455
3.6564 | -83.4555
3.6567 | -83.4562
3.657 | -83.4568
3.6574 | -83.4579
3.6577 | -83.458
3.658 | -83.4592
3.6583 | -83.461
3.6586 | -83.4616
3.6589 | -83.4621
3.6592 | -83.4627
3.6595 | -83.4644
3.6598 | -83.4661
3.6601 | -83.4669
3.6604 | -83.4678
3.6608 | -83.4691
3.6611 | -83.4685
3.6614 | -83.4678
3.6617 | -83.4683
3.662 | -83.4704
3.6623 | -83.4713
3.6626 | -83.4719
3.6629 | -83.4731
3.6632 | -83.4743
3.6635 | -83.4752
3.6638 | -83.4771
3.6641 | -83.4785
3.6645 | -83.4801
3.6648 | -83.4825
3.6651 | -83.4841
3.6654 | -83.4858
3.6657 | -83.487
3.666 | -83.4885
3.6663 | -83.4917
3.6666 | -83.4957
3.6669 | -83.4981
3.6672 | -83.4992
3.6675 | -83.4999
3.6679 | -83.5015
3.6682 | -83.5029
3.6685 | -83.5036
3.6688 | -83.5047
3.6691 | -83.5057
3.6694 | -83.5067
3.6697 | -83.5096
3.67 | -83.512
3.6703 | -83.5144
3.6706 | -83.5157
3.6709 | -83.5164
3.6712 | -83.5169
3.6716 | -83.5174
3.6719 | -83.5185
3.6722 | -83.5198
3.6725 | -83.5213
3.6728 | -83.5229
3.6731 | -83.525
3.6734 | -83.5265
3.6737 | -83.5301
3.674 | -83.5317
3.6743 | -83.5332
3.6746 | -83.5335
3.675 | -83.5355
3.6753 | -83.5372
3.6756 | -83.5389
3.6759 | -83.5411
3.6762 | -83.5423
3.6765 | -83.5444
3.6768 | -83.5472
3.6771 | -83.55
3.6774 | -83.5528
3.6777 | -83.5542
3.678 | -83.556
3.6784 | -83.5565
3.6787 | -83.5569
3.679 | -83.5588
3.6793 | -83.5607
3.6796 | -83.5625
3.6799 | -83.5625
3.6802 | -83.5642
3.6805 | -83.5665
3.6808 | -83.5681
3.6811 | -83.57
3.6814 | -83.5726
3.6817 | -83.5743
3.6821 | -83.5752
3.6824 | -83.5773
3.6827 | -83.5802
3.683 | -83.5823
3.6833 | -83.5834
3.6836 | -83.5851
3.6839 | -83.587
3.6842 | -83.5889
3.6845 | -83.5906
3.6848 | -83.592
3.6851 | -83.5938
3.6855 | -83.5955
3.6858 | -83.597
3.6861 | -83.6
3.6864 | -83.6025
3.6867 | -83.6048
3.687 | -83.6081
3.6873 | -83.6108
3.6876 | -83.6127
3.6879 | -83.6155
3.6882 | -83.6173
3.6885 | -83.6192
3.6888 | -83.6212
3.6892 | -83.6221
3.6895 | -83.6245
3.6898 | -83.6266
3.6901 | -83.6288
3.6904 | -83.6313
3.6907 | -83.6331
3.691 | -83.636
3.6913 | -83.6376
3.6916 | -83.6391
3.6919 | -83.6413
3.6922 | -83.6433
3.6926 | -83.6446
3.6929 | -83.6452
3.6932 | -83.6469
3.6935 | -83.6489
3.6938 | -83.6515
3.6941 | -83.6557
3.6944 | -83.6594
3.6947 | -83.6632
3.695 | -83.6658
3.6953 | -83.6692
3.6956 | -83.6726
3.6959 | -83.6742
3.6963 | -83.6757
3.6966 | -83.678
3.6969 | -83.6798
3.6972 | -83.682
3.6975 | -83.6836
3.6978 | -83.687
3.6981 | -83.6901
3.6984 | -83.6919
3.6987 | -83.6938
3.699 | -83.6961
3.6993 | -83.6994
3.6997 | -83.7014
3.7 | -83.7029
3.7003 | -83.7045
3.7006 | -83.7059
3.7009 | -83.7085
3.7012 | -83.7106
3.7015 | -83.7123
3.7018 | -83.7138
3.7021 | -83.7163
3.7024 | -83.7185
3.7027 | -83.7205
3.7031 | -83.7228
3.7034 | -83.7244
3.7037 | -83.726
3.704 | -83.7288
3.7043 | -83.7318
3.7046 | -83.7338
3.7049 | -83.7359
3.7052 | -83.7384
3.7055 | -83.7431
3.7058 | -83.7426
3.7061 | -83.7449
3.7064 | -83.7471
3.7068 | -83.7493
3.7071 | -83.7514
3.7074 | -83.7533
3.7077 | -83.7549
3.708 | -83.7557
3.7083 | -83.7578
3.7086 | -83.7599
3.7089 | -83.762
3.7092 | -83.7639
3.7095 | -83.765
3.7098 | -83.767
3.7102 | -83.7707
3.7105 | -83.7767
3.7108 | -83.7812
3.7111 | -83.7822
3.7114 | -83.783
3.7117 | -83.7851
3.712 | -83.787
3.7123 | -83.7888
3.7126 | -83.7902
3.7129 | -83.792
3.7132 | -83.794
3.7136 | -83.7954
3.7139 | -83.7974
3.7142 | -83.7995
3.7145 | -83.8022
3.7148 | -83.8036
3.7151 | -83.8076
3.7154 | -83.8107
3.7157 | -83.8132
3.716 | -83.8154
3.7163 | -83.8175
3.7166 | -83.8205
3.7169 | -83.8225
3.7173 | -83.8243
3.7176 | -83.8264
3.7179 | -83.8287
3.7182 | -83.8307
3.7185 | -83.8318
3.7188 | -83.8332
3.7191 | -83.8357
3.7194 | -83.8401
3.7197 | -83.8439
3.72 | -83.8472
3.7203 | -83.8494
3.7207 | -83.8489
3.721 | -83.8494
3.7213 | -83.8506
3.7216 | -83.8545
3.7219 | -83.8592
3.7222 | -83.8623
3.7225 | -83.8644
3.7228 | -83.868
3.7231 | -83.8705
3.7234 | -83.8724
3.7237 | -83.8749
3.724 | -83.877
3.7244 | -83.8793
3.7247 | -83.8818
3.725 | -83.8831
3.7253 | -83.8848
3.7256 | -83.8878
3.7259 | -83.8914
3.7262 | -83.8925
3.7265 | -83.8956
3.7268 | -83.8982
3.7271 | -83.9007
3.7274 | -83.9026
3.7278 | -83.9047
3.7281 | -83.9069
3.7284 | -83.9082
3.7287 | -83.9103
3.729 | -83.9131
3.7293 | -83.9151
3.7296 | -83.917
3.7299 | -83.918
3.7302 | -83.92
3.7305 | -83.922
3.7308 | -83.924
3.7311 | -83.9261
3.7315 | -83.9285
3.7318 | -83.9303
3.7321 | -83.9325
3.7324 | -83.9349
3.7327 | -83.9365
3.733 | -83.9379
3.7333 | -83.9395
3.7336 | -83.9414
3.7339 | -83.9429
3.7342 | -83.9451
3.7345 | -83.9474
3.7349 | -83.9498
3.7352 | -83.9528
3.7355 | -83.9542
3.7358 | -83.9565
3.7361 | -83.9593
3.7364 | -83.9616
3.7367 | -83.9639
3.737 | -83.966
3.7373 | -83.9675
3.7376 | -83.9694
3.7379 | -83.971
3.7383 | -83.9729
3.7386 | -83.9745
3.7389 | -83.9764
3.7392 | -83.9784
3.7395 | -83.9807
3.7398 | -83.9838
3.7401 | -83.9879
3.7404 | -83.9919
3.7407 | -83.9941
3.741 | -83.9964
3.7413 | -83.9987
3.7416 | -84.0011
3.742 | -84.0033
3.7423 | -84.0055
3.7426 | -84.0077
3.7429 | -84.0099
3.7432 | -84.0122
3.7435 | -84.014
3.7438 | -84.0167
3.7441 | -84.0186
3.7444 | -84.0208
3.7447 | -84.0246
3.745 | -84.0289
3.7454 | -84.0314
3.7457 | -84.0319
3.746 | -84.032
3.7463 | -84.0338
3.7466 | -84.0358
3.7469 | -84.0398
3.7472 | -84.0415
3.7475 | -84.0405
3.7478 | -84.0402
3.7481 | -84.0405
3.7484 | -84.0426
3.7487 | -84.0455
3.7491 | -84.0479
3.7494 | -84.0491
3.7497 | -84.0516
3.75 | -84.0561
3.7503 | -84.0584
3.7506 | -84.0605
3.7509 | -84.0619
3.7512 | -84.064
3.7515 | -84.0673
3.7518 | -84.0685
3.7521 | -84.0687
3.7525 | -84.0687
3.7528 | -84.0698
3.7531 | -84.0707
3.7534 | -84.072
3.7537 | -84.0738
3.754 | -84.0781
3.7543 | -84.0849
3.7546 | -84.0872
3.7549 | -84.0888
3.7552 | -84.0903
3.7555 | -84.092
3.7559 | -84.0946
3.7562 | -84.0964
3.7565 | -84.0987
3.7568 | -84.1004
3.7571 | -84.1027
3.7574 | -84.1046
3.7577 | -84.1071
3.758 | -84.1104
3.7583 | -84.1117
3.7586 | -84.1126
3.7589 | -84.1117
3.7592 | -84.1124
3.7596 | -84.115
3.7599 | -84.1169
3.7602 | -84.1191
3.7605 | -84.1206
3.7608 | -84.1211
3.7611 | -84.1221
3.7614 | -84.124
3.7617 | -84.1262
3.762 | -84.1287
3.7623 | -84.1301
3.7626 | -84.1321
3.763 | -84.1318
3.7633 | -84.1301
3.7636 | -84.1303
3.7639 | -84.1315
3.7642 | -84.1359
3.7645 | -84.1398
3.7648 | -84.1417
3.7651 | -84.1431
3.7654 | -84.1446
3.7657 | -84.146
3.766 | -84.1479
3.7663 | -84.1497
3.7667 | -84.1514
3.767 | -84.1536
3.7673 | -84.1549
3.7676 | -84.1572
3.7679 | -84.1596
3.7682 | -84.1608
3.7685 | -84.1625
3.7688 | -84.1634
3.7691 | -84.1647
3.7694 | -84.1648
3.7697 | -84.1654
3.7701 | -84.1668
3.7704 | -84.1682
3.7707 | -84.1707
3.771 | -84.1732
3.7713 | -84.1748
3.7716 | -84.1765
3.7719 | -84.1781
3.7722 | -84.1789
3.7725 | -84.1811
3.7728 | -84.1825
3.7731 | -84.1822
3.7735 | -84.1827
3.7738 | -84.1838
3.7741 | -84.1852
3.7744 | -84.1864
3.7747 | -84.1867
3.775 | -84.1866
3.7753 | -84.1874
3.7756 | -84.1886
3.7759 | -84.1891
3.7762 | -84.1899
3.7765 | -84.1913
3.7768 | -84.1927
3.7772 | -84.1912
3.7775 | -84.1924
3.7778 | -84.1934
3.7781 | -84.1946
3.7784 | -84.1963
3.7787 | -84.1979
3.779 | -84.199
3.7793 | -84.1998
3.7796 | -84.2003
3.7799 | -84.2025
3.7802 | -84.2052
3.7806 | -84.2065
3.7809 | -84.2073
3.7812 | -84.2084
3.7815 | -84.2085
3.7818 | -84.2102
3.7821 | -84.2123
3.7824 | -84.2132
3.7827 | -84.2147
3.783 | -84.2169
3.7833 | -84.2185
3.7836 | -84.2199
3.7839 | -84.2208
3.7843 | -84.2217
3.7846 | -84.2225
3.7849 | -84.2227
3.7852 | -84.224
3.7855 | -84.2253
3.7858 | -84.2263
3.7861 | -84.2251
3.7864 | -84.2253
3.7867 | -84.2253
3.787 | -84.2298
3.7873 | -84.2358
3.7877 | -84.2374
3.788 | -84.2383
3.7883 | -84.2405
3.7886 | -84.2415
3.7889 | -84.242
3.7892 | -84.2426
3.7895 | -84.2428
3.7898 | -84.2428
3.7901 | -84.2439
3.7904 | -84.245
3.7907 | -84.2457
3.7911 | -84.2461
3.7914 | -84.2478
3.7917 | -84.2498
3.792 | -84.2515
3.7923 | -84.2531
3.7926 | -84.2549
3.7929 | -84.2567
3.7932 | -84.2579
3.7935 | -84.2586
3.7938 | -84.2574
3.7941 | -84.257
3.7944 | -84.2565
3.7948 | -84.2565
3.7951 | -84.2563
3.7954 | -84.2548
3.7957 | -84.2535
3.796 | -84.2524
3.7963 | -84.251
3.7966 | -84.2508
3.7969 | -84.2516
3.7972 | -84.2522
3.7975 | -84.2524
3.7978 | -84.2529
3.7982 | -84.2536
3.7985 | -84.254
3.7988 | -84.2552
3.7991 | -84.2565
3.7994 | -84.256
3.7997 | -84.2551
3.8 | -84.2526
3.8003 | -84.2499
3.8006 | -84.248
3.8009 | -84.2476
3.8012 | -84.2478
3.8015 | -84.2472
3.8019 | -84.2482
3.8022 | -84.2485
3.8025 | -84.2481
3.8028 | -84.2488
3.8031 | -84.25
3.8034 | -84.2498
3.8037 | -84.2476
3.804 | -84.247
3.8043 | -84.2483
3.8046 | -84.2519
3.8049 | -84.2567
3.8053 | -84.2596
3.8056 | -84.26
3.8059 | -84.2608
3.8062 | -84.2601
3.8065 | -84.2593
3.8068 | -84.26
3.8071 | -84.2602
3.8074 | -84.2602
3.8077 | -84.2598
3.808 | -84.2596
3.8083 | -84.2593
3.8087 | -84.2598
3.809 | -84.2604
3.8093 | -84.2602
3.8096 | -84.2602
3.8099 | -84.2604
3.8102 | -84.2604
3.8105 | -84.2599
3.8108 | -84.2583
3.8111 | -84.2569
3.8114 | -84.257
3.8117 | -84.2562
3.812 | -84.2551
3.8124 | -84.2551
3.8127 | -84.2544
3.813 | -84.2537
3.8133 | -84.2538
3.8136 | -84.2541
3.8139 | -84.2547
3.8142 | -84.2552
3.8145 | -84.2545
3.8148 | -84.2535
3.8151 | -84.2523
3.8154 | -84.2504
3.8158 | -84.25
3.8161 | -84.2502
3.8164 | -84.2495
3.8167 | -84.2497
3.817 | -84.2505
3.8173 | -84.2527
3.8176 | -84.2552
3.8179 | -84.2565
3.8182 | -84.2567
3.8185 | -84.2565
3.8188 | -84.2559
3.8191 | -84.2553
3.8195 | -84.2547
3.8198 | -84.2543
3.8201 | -84.2542
3.8204 | -84.2549
3.8207 | -84.2548
3.821 | -84.2543
3.8213 | -84.2541
3.8216 | -84.2539
3.8219 | -84.253
3.8222 | -84.2523
3.8225 | -84.2519
3.8229 | -84.2513
3.8232 | -84.2503
3.8235 | -84.2507
3.8238 | -84.2509
3.8241 | -84.2509
3.8244 | -84.252
3.8247 | -84.2513
3.825 | -84.2506
3.8253 | -84.2499
3.8256 | -84.2505
3.8259 | -84.2496
3.8262 | -84.2486
3.8266 | -84.2478
3.8269 | -84.2477
3.8272 | -84.247
3.8275 | -84.2462
3.8278 | -84.2471
3.8281 | -84.2486
3.8284 | -84.2481
3.8287 | -84.2475
3.829 | -84.2467
3.8293 | -84.2445
3.8296 | -84.2421
3.83 | -84.2382
3.8303 | -84.2349
3.8306 | -84.2332
3.8309 | -84.2327
3.8312 | -84.2318
3.8315 | -84.2294
3.8318 | -84.2286
3.8321 | -84.2272
3.8324 | -84.2276
3.8327 | -84.2285
3.833 | -84.2284
3.8334 | -84.2289
3.8337 | -84.2288
3.834 | -84.2295
3.8343 | -84.228
3.8346 | -84.2268
3.8349 | -84.2269
3.8352 | -84.2257
3.8355 | -84.2242
3.8358 | -84.2231
3.8361 | -84.2223
3.8364 | -84.2204
3.8367 | -84.2187
3.8371 | -84.2165
3.8374 | -84.2147
3.8377 | -84.2135
3.838 | -84.212
3.8383 | -84.2093
3.8386 | -84.2077
3.8389 | -84.2064
3.8392 | -84.2048
3.8395 | -84.204
3.8398 | -84.2019
3.8401 | -84.2
3.8405 | -84.1983
3.8408 | -84.197
3.8411 | -84.196
3.8414 | -84.1949
3.8417 | -84.1945
3.842 | -84.1935
3.8423 | -84.1916
3.8426 | -84.1901
3.8429 | -84.1886
3.8432 | -84.187
3.8435 | -84.1865
3.8439 | -84.1844
3.8442 | -84.1818
3.8445 | -84.1801
3.8448 | -84.1793
3.8451 | -84.1781
3.8454 | -84.1768
3.8457 | -84.1761
3.846 | -84.1751
3.8463 | -84.174
3.8466 | -84.1732
3.8469 | -84.1715
3.8472 | -84.1695
3.8476 | -84.1663
3.8479 | -84.1638
3.8482 | -84.1626
3.8485 | -84.1616
3.8488 | -84.1601
3.8491 | -84.1606
3.8494 | -84.1599
3.8497 | -84.1586
3.85 | -84.157
3.8503 | -84.1557
3.8506 | -84.1547
3.851 | -84.154
3.8513 | -84.1526
3.8516 | -84.1509
3.8519 | -84.1501
3.8522 | -84.1497
3.8525 | -84.15
3.8528 | -84.1512
3.8531 | -84.1508
3.8534 | -84.1493
3.8537 | -84.1484
3.854 | -84.1471
3.8543 | -84.1456
3.8547 | -84.1446
3.855 | -84.1433
3.8553 | -84.1415
3.8556 | -84.1403
3.8559 | -84.138
3.8562 | -84.1361
3.8565 | -84.1349
3.8568 | -84.1344
3.8571 | -84.1332
3.8574 | -84.1326
3.8577 | -84.1311
3.8581 | -84.1297
3.8584 | -84.1284
3.8587 | -84.1269
3.859 | -84.1258
3.8593 | -84.1247
3.8596 | -84.123
3.8599 | -84.1215
3.8602 | -84.1202
3.8605 | -84.1188
3.8608 | -84.1173
3.8611 | -84.1161
3.8615 | -84.1154
3.8618 | -84.1144
3.8621 | -84.1133
3.8624 | -84.1102
3.8627 | -84.1054
3.863 | -84.1027
3.8633 | -84.0996
3.8636 | -84.0992
3.8639 | -84.099
3.8642 | -84.0968
3.8645 | -84.0954
3.8648 | -84.094
3.8652 | -84.0924
3.8655 | -84.0911
3.8658 | -84.0898
3.8661 | -84.0887
3.8664 | -84.0863
3.8667 | -84.0843
3.867 | -84.0824
3.8673 | -84.0821
3.8676 | -84.0809
3.8679 | -84.0796
3.8682 | -84.0783
3.8686 | -84.0754
3.8689 | -84.0729
3.8692 | -84.0723
3.8695 | -84.0716
3.8698 | -84.0713
3.8701 | -84.0683
3.8704 | -84.0654
3.8707 | -84.0639
3.871 | -84.0624
3.8713 | -84.0608
3.8716 | -84.0597
3.8719 | -84.0588
3.8723 | -84.0577
3.8726 | -84.0568
3.8729 | -84.0554
3.8732 | -84.0537
3.8735 | -84.0516
3.8738 | -84.0493
3.8741 | -84.0476
3.8744 | -84.0465
3.8747 | -84.0454
3.875 | -84.0446
3.8753 | -84.0432
3.8757 | -84.0408
3.876 | -84.0379
3.8763 | -84.0376
3.8766 | -84.0367
3.8769 | -84.0361
3.8772 | -84.036
3.8775 | -84.035
3.8778 | -84.0329
3.8781 | -84.0314
3.8784 | -84.0292
3.8787 | -84.0269
3.879 | -84.0251
3.8794 | -84.0229
3.8797 | -84.0206
3.88 | -84.0191
3.8803 | -84.0151
3.8806 | -84.0115
3.8809 | -84.0096
3.8812 | -84.0084
3.8815 | -84.0057
3.8818 | -84.0028
3.8821 | -84.0018
3.8824 | -84.0011
3.8828 | -83.9997
3.8831 | -83.9973
3.8834 | -83.9958
3.8837 | -83.9941
3.884 | -83.9929
3.8843 | -83.9905
3.8846 | -83.9888
3.8849 | -83.9873
3.8852 | -83.9859
3.8855 | -83.9844
3.8858 | -83.9833
3.8862 | -83.9823
3.8865 | -83.9814
3.8868 | -83.9803
3.8871 | -83.9791
3.8874 | -83.978
3.8877 | -83.9777
3.888 | -83.976
3.8883 | -83.9747
3.8886 | -83.973
3.8889 | -83.9724
3.8892 | -83.9703
3.8895 | -83.9681
3.8899 | -83.9676
3.8902 | -83.9662
3.8905 | -83.9651
3.8908 | -83.9632
3.8911 | -83.9617
3.8914 | -83.9602
3.8917 | -83.9589
3.892 | -83.9577
3.8923 | -83.9555
3.8926 | -83.9534
3.8929 | -83.9524
3.8933 | -83.95
3.8936 | -83.9478
3.8939 | -83.946
3.8942 | -83.9443
3.8945 | -83.9428
3.8948 | -83.941
3.8951 | -83.9395
3.8954 | -83.9383
3.8957 | -83.9371
3.896 | -83.9357
3.8963 | -83.9334
3.8966 | -83.9299
3.897 | -83.9266
3.8973 | -83.926
3.8976 | -83.9258
3.8979 | -83.9243
3.8982 | -83.9224
3.8985 | -83.9211
3.8988 | -83.9198
3.8991 | -83.9184
3.8994 | -83.9173
3.8997 | -83.9156
3.9 | -83.9149
3.9004 | -83.914
3.9007 | -83.9129
3.901 | -83.9115
3.9013 | -83.9109
3.9016 | -83.9093
3.9019 | -83.9081
3.9022 | -83.9061
3.9025 | -83.9049
3.9028 | -83.9017
3.9031 | -83.9014
3.9034 | -83.8999
3.9038 | -83.8984
3.9041 | -83.896
3.9044 | -83.8926
3.9047 | -83.8915
3.905 | -83.8907
3.9053 | -83.8889
3.9056 | -83.8871
3.9059 | -83.8862
3.9062 | -83.885
3.9065 | -83.8835
3.9068 | -83.8819
3.9071 | -83.8801
3.9075 | -83.8794
3.9078 | -83.8789
3.9081 | -83.8774
3.9084 | -83.8757
3.9087 | -83.8752
3.909 | -83.8757
3.9093 | -83.8744
3.9096 | -83.8727
3.9099 | -83.8704
3.9102 | -83.8687
3.9105 | -83.8674
3.9109 | -83.8678
3.9112 | -83.8664
3.9115 | -83.8652
3.9118 | -83.864
3.9121 | -83.8628
3.9124 | -83.8625
3.9127 | -83.8613
3.913 | -83.86
3.9133 | -83.8589
3.9136 | -83.8577
3.9139 | -83.8556
3.9142 | -83.8525
3.9146 | -83.8498
3.9149 | -83.8484
3.9152 | -83.8468
3.9155 | -83.8462
3.9158 | -83.8446
3.9161 | -83.8429
3.9164 | -83.8429
3.9167 | -83.8416
3.917 | -83.8406
3.9173 | -83.8398
3.9176 | -83.8388
3.918 | -83.8371
3.9183 | -83.8348
3.9186 | -83.8329
3.9189 | -83.8321
3.9192 | -83.832
3.9195 | -83.8309
3.9198 | -83.8303
3.9201 | -83.8296
3.9204 | -83.8285
3.9207 | -83.8272
3.921 | -83.8269
3.9214 | -83.8247
3.9217 | -83.8213
3.922 | -83.8193
3.9223 | -83.818
3.9226 | -83.8168
3.9229 | -83.8154
3.9232 | -83.8145
3.9235 | -83.814
3.9238 | -83.8121
3.9241 | -83.8101
3.9244 | -83.809
3.9247 | -83.8088
3.9251 | -83.8077
3.9254 | -83.8071
3.9257 | -83.8053
3.926 | -83.804
3.9263 | -83.8032
3.9266 | -83.8024
3.9269 | -83.8013
3.9272 | -83.7992
3.9275 | -83.7967
3.9278 | -83.7954
3.9281 | -83.7943
3.9285 | -83.7944
3.9288 | -83.7927
3.9291 | -83.7906
3.9294 | -83.7885
3.9297 | -83.7871
3.93 | -83.7855
3.9303 | -83.7847
3.9306 | -83.7836
3.9309 | -83.7826
3.9312 | -83.7819
3.9315 | -83.7811
3.9318 | -83.7797
3.9322 | -83.7787
3.9325 | -83.7792
3.9328 | -83.779
3.9331 | -83.778
3.9334 | -83.7771
3.9337 | -83.7761
3.934 | -83.7753
3.9343 | -83.7744
3.9346 | -83.7736
3.9349 | -83.7728
3.9352 | -83.7723
3.9356 | -83.7719
3.9359 | -83.7717
3.9362 | -83.7708
3.9365 | -83.7699
3.9368 | -83.7689
3.9371 | -83.767
3.9374 | -83.7669
3.9377 | -83.7661
3.938 | -83.7653
3.9383 | -83.7649
3.9386 | -83.7643
3.939 | -83.7637
3.9393 | -83.7646
3.9396 | -83.7643
3.9399 | -83.7641
3.9402 | -83.7641
3.9405 | -83.7626
3.9408 | -83.7611
3.9411 | -83.7601
3.9414 | -83.759
3.9417 | -83.7584
3.942 | -83.7581
3.9423 | -83.7576
3.9427 | -83.757
3.943 | -83.7574
3.9433 | -83.7582
3.9436 | -83.7588
3.9439 | -83.7582
3.9442 | -83.7581
3.9445 | -83.7577
3.9448 | -83.7588
3.9451 | -83.7586
3.9454 | -83.758
3.9457 | -83.7573
3.9461 | -83.757
3.9464 | -83.7577
3.9467 | -83.7575
3.947 | -83.7569
3.9473 | -83.7563
3.9476 | -83.7549
3.9479 | -83.753
3.9482 | -83.7519
3.9485 | -83.7512
3.9488 | -83.7505
3.9491 | -83.7501
3.9494 | -83.7492
3.9498 | -83.7486
3.9501 | -83.7479
3.9504 | -83.7475
3.9507 | -83.7469
3.951 | -83.7446
3.9513 | -83.7424
3.9516 | -83.7418
3.9519 | -83.7404
3.9522 | -83.739
3.9525 | -83.7384
3.9528 | -83.738
3.9532 | -83.7379
3.9535 | -83.7375
3.9538 | -83.7372
3.9541 | -83.7371
3.9544 | -83.7379
3.9547 | -83.7382
3.955 | -83.7378
3.9553 | -83.7385
3.9556 | -83.7381
3.9559 | -83.7369
3.9562 | -83.7366
3.9566 | -83.7361
3.9569 | -83.7352
3.9572 | -83.735
3.9575 | -83.7342
3.9578 | -83.733
3.9581 | -83.7316
3.9584 | -83.7311
3.9587 | -83.7318
3.959 | -83.7318
3.9593 | -83.7308
3.9596 | -83.7301
3.9599 | -83.7297
3.9603 | -83.7291
3.9606 | -83.7286
3.9609 | -83.7301
3.9612 | -83.732
3.9615 | -83.7321
3.9618 | -83.7321
3.9621 | -83.7321
3.9624 | -83.7317
3.9627 | -83.7314
3.963 | -83.7319
3.9633 | -83.7312
3.9637 | -83.73
3.964 | -83.7296
3.9643 | -83.7307
3.9646 | -83.7308
3.9649 | -83.7301
3.9652 | -83.7289
3.9655 | -83.7285
3.9658 | -83.7281
3.9661 | -83.7275
3.9664 | -83.7277
3.9667 | -83.7279
3.967 | -83.7278
3.9674 | -83.7267
3.9677 | -83.7247
3.968 | -83.724
3.9683 | -83.7242
3.9686 | -83.7236
3.9689 | -83.7227
3.9692 | -83.7195
3.9695 | -83.7155
3.9698 | -83.7149
3.9701 | -83.7151
3.9704 | -83.7145
3.9708 | -83.7137
3.9711 | -83.7133
3.9714 | -83.7122
3.9717 | -83.7109
3.972 | -83.7105
3.9723 | -83.7109
3.9726 | -83.7135
3.9729 | -83.7147
3.9732 | -83.7145
3.9735 | -83.7132
3.9738 | -83.7118
3.9741 | -83.7108
3.9745 | -83.7099
3.9748 | -83.7097
3.9751 | -83.7106
3.9754 | -83.7104
3.9757 | -83.7099
3.976 | -83.7092
3.9763 | -83.7095
3.9766 | -83.7102
3.9769 | -83.7097
3.9772 | -83.7093
3.9775 | -83.709
3.9779 | -83.7088
3.9782 | -83.7097
3.9785 | -83.711
3.9788 | -83.7111
3.9791 | -83.7109
3.9794 | -83.7107
3.9797 | -83.711
3.98 | -83.7109
3.9803 | -83.7106
3.9806 | -83.7109
3.9809 | -83.711
3.9813 | -83.7098
3.9816 | -83.7095
3.9819 | -83.7097
3.9822 | -83.7099
3.9825 | -83.7099
3.9828 | -83.71
3.9831 | -83.7099
3.9834 | -83.7094
3.9837 | -83.7093
3.984 | -83.7086
3.9843 | -83.7084
3.9846 | -83.7087
3.985 | -83.7089
3.9853 | -83.7086
3.9856 | -83.7077
3.9859 | -83.7065
3.9862 | -83.707
3.9865 | -83.7073
3.9868 | -83.7071
3.9871 | -83.707
3.9874 | -83.707
3.9877 | -83.7069
3.988 | -83.7072
3.9884 | -83.7095
3.9887 | -83.7108
3.989 | -83.7107
3.9893 | -83.7098
3.9896 | -83.7097
3.9899 | -83.7088
3.9902 | -83.7088
3.9905 | -83.7079
3.9908 | -83.7069
3.9911 | -83.706
3.9914 | -83.7058
3.9918 | -83.7056
3.9921 | -83.7058
3.9924 | -83.7061
3.9927 | -83.7058
3.993 | -83.7055
3.9933 | -83.7067
3.9936 | -83.7061
3.9939 | -83.7061
3.9942 | -83.7068
3.9945 | -83.7075
3.9948 | -83.7074
3.9951 | -83.7073
3.9955 | -83.7063
3.9958 | -83.7055
3.9961 | -83.7052
3.9964 | -83.705
3.9967 | -83.7049
3.997 | -83.7045
3.9973 | -83.7039
3.9976 | -83.7035
3.9979 | -83.7038
3.9982 | -83.7045
3.9985 | -83.705
3.9989 | -83.7043
3.9992 | -83.7042
3.9995 | -83.7041
3.9998 | -83.704
4.0001 | -83.7026
4.0004 | -83.7022
4.0007 | -83.7026
4.001 | -83.7029
4.0013 | -83.7031
4.0016 | -83.7029
4.0019 | -83.7024
4.0022 | -83.7028
4.0026 | -83.7031
4.0029 | -83.7034
4.0032 | -83.703
4.0035 | -83.703
4.0038 | -83.7027
4.0041 | -83.7017
4.0044 | -83.702
4.0047 | -83.702
4.005 | -83.7022
4.0053 | -83.7019
4.0056 | -83.701
4.006 | -83.7005
4.0063 | -83.701
4.0066 | -83.7013
4.0069 | -83.7018
4.0072 | -83.7013
4.0075 | -83.7018
4.0078 | -83.7024
4.0081 | -83.7026
4.0084 | -83.7017
4.0087 | -83.7015
4.009 | -83.7
4.0093 | -83.6994
4.0097 | -83.6979
4.01 | -83.6968
4.0103 | -83.6967
4.0106 | -83.6966
4.0109 | -83.6958
4.0112 | -83.6955
4.0115 | -83.6955
4.0118 | -83.6954
4.0121 | -83.6951
4.0124 | -83.6952
4.0127 | -83.6955
4.0131 | -83.6953
4.0134 | -83.6952
4.0137 | -83.6955
4.014 | -83.6953
4.0143 | -83.6949
4.0146 | -83.6946
4.0149 | -83.6945
4.0152 | -83.696
4.0155 | -83.6973
4.0158 | -83.6964
4.0161 | -83.6963
4.0165 | -83.696
4.0168 | -83.6963
4.0171 | -83.6987
4.0174 | -83.7007
4.0177 | -83.7008
4.018 | -83.7023
4.0183 | -83.7022
4.0186 | -83.7014
4.0189 | -83.7012
4.0192 | -83.7019
4.0195 | -83.7026
4.0198 | -83.7022
4.0202 | -83.7018
4.0205 | -83.7019
4.0208 | -83.7022
4.0211 | -83.7027
4.0214 | -83.7035
4.0217 | -83.7034
4.022 | -83.7035
4.0223 | -83.7039
4.0226 | -83.7044
4.0229 | -83.7049
4.0232 | -83.7058
4.0236 | -83.7065
4.0239 | -83.7066
4.0242 | -83.7069
4.0245 | -83.7066
4.0248 | -83.7066
4.0251 | -83.7067
4.0254 | -83.707
4.0257 | -83.7072
4.026 | -83.7072
4.0263 | -83.7075
4.0266 | -83.7085
4.0269 | -83.7077
4.0273 | -83.7066
4.0276 | -83.7064
4.0279 | -83.7065
4.0282 | -83.7051
4.0285 | -83.7035
4.0288 | -83.7032
4.0291 | -83.7035
4.0294 | -83.7053
4.0297 | -83.7064
4.03 | -83.7064
4.0303 | -83.7061
4.0307 | -83.7057
4.031 | -83.7055
4.0313 | -83.7058
4.0316 | -83.707
4.0319 | -83.7079
4.0322 | -83.708
4.0325 | -83.7083
4.0328 | -83.7081
4.0331 | -83.7084
4.0334 | -83.7085
4.0337 | -83.709
4.0341 | -83.7107
4.0344 | -83.7119
4.0347 | -83.7118
4.035 | -83.712
4.0353 | -83.7122
4.0356 | -83.7135
4.0359 | -83.7136
4.0362 | -83.714
4.0365 | -83.7144
4.0368 | -83.7149
4.0371 | -83.7148
4.0374 | -83.7151
4.0378 | -83.7156
4.0381 | -83.7154
4.0384 | -83.7141
4.0387 | -83.7134
4.039 | -83.7133
4.0393 | -83.7135
4.0396 | -83.7142
4.0399 | -83.7145
4.0402 | -83.7154
4.0405 | -83.7154
4.0408 | -83.7141
4.0412 | -83.7141
4.0415 | -83.7142
4.0418 | -83.7149
4.0421 | -83.7153
4.0424 | -83.715
4.0427 | -83.715
4.043 | -83.7148
4.0433 | -83.7148
4.0436 | -83.715
4.0439 | -83.7145
4.0442 | -83.7156
4.0445 | -83.7154
4.0449 | -83.7171
4.0452 | -83.7193
4.0455 | -83.7201
4.0458 | -83.7207
4.0461 | -83.7216
4.0464 | -83.7234
4.0467 | -83.7252
4.047 | -83.7254
4.0473 | -83.7253
4.0476 | -83.7264
4.0479 | -83.7273
4.0483 | -83.7273
4.0486 | -83.7277
4.0489 | -83.7298
4.0492 | -83.7307
4.0495 | -83.7303
4.0498 | -83.7301
4.0501 | -83.7303
4.0504 | -83.7306
4.0507 | -83.7309
4.051 | -83.7303
4.0513 | -83.7311
4.0517 | -83.7315
4.052 | -83.73
4.0523 | -83.7285
4.0526 | -83.7294
4.0529 | -83.7302
4.0532 | -83.7318
4.0535 | -83.7321
4.0538 | -83.7325
4.0541 | -83.7336
4.0544 | -83.7335
4.0547 | -83.7327
4.055 | -83.7326
4.0554 | -83.733
4.0557 | -83.7339
4.056 | -83.7346
4.0563 | -83.7356
4.0566 | -83.7367
4.0569 | -83.7376
4.0572 | -83.7383
4.0575 | -83.7386
4.0578 | -83.7384
4.0581 | -83.7389
4.0584 | -83.7395
4.0588 | -83.7409
4.0591 | -83.7422
4.0594 | -83.7436
4.0597 | -83.7442
4.06 | -83.7442
4.0603 | -83.7447
4.0606 | -83.7455
4.0609 | -83.746
4.0612 | -83.7466
4.0615 | -83.7469
4.0618 | -83.7486
4.0621 | -83.75
4.0625 | -83.7528
4.0628 | -83.7538
4.0631 | -83.7535
4.0634 | -83.7528
4.0637 | -83.7521
4.064 | -83.7528
4.0643 | -83.753
4.0646 | -83.7532
4.0649 | -83.7537
4.0652 | -83.7543
4.0655 | -83.7547
4.0659 | -83.7548
4.0662 | -83.7553
4.0665 | -83.7568
4.0668 | -83.7593
4.0671 | -83.7606
4.0674 | -83.7618
4.0677 | -83.7636
4.068 | -83.7646
4.0683 | -83.7655
4.0686 | -83.7679
4.0689 | -83.77
4.0693 | -83.7719
4.0696 | -83.7743
4.0699 | -83.7763
4.0702 | -83.7777
4.0705 | -83.7779
4.0708 | -83.7776
4.0711 | -83.7793
4.0714 | -83.7799
4.0717 | -83.7807
4.072 | -83.7816
4.0723 | -83.7823
4.0726 | -83.7818
4.073 | -83.7823
4.0733 | -83.7843
4.0736 | -83.7876
4.0739 | -83.7913
4.0742 | -83.7934
4.0745 | -83.7945
4.0748 | -83.796
4.0751 | -83.7972
4.0754 | -83.7974
4.0757 | -83.7984
4.076 | -83.7992
4.0764 | -83.7997
4.0767 | -83.7985
4.077 | -83.7992
4.0773 | -83.8004
4.0776 | -83.802
4.0779 | -83.803
4.0782 | -83.8036
4.0785 | -83.8047
4.0788 | -83.8053
4.0791 | -83.8059
4.0794 | -83.8065
4.0797 | -83.806
4.0801 | -83.807
4.0804 | -83.8084
4.0807 | -83.8096
4.081 | -83.8101
4.0813 | -83.811
4.0816 | -83.8125
4.0819 | -83.8141
4.0822 | -83.8153
4.0825 | -83.8161
4.0828 | -83.8171
4.0831 | -83.8176
4.0835 | -83.8183
4.0838 | -83.8192
4.0841 | -83.8201
4.0844 | -83.8212
4.0847 | -83.8224
4.085 | -83.8243
4.0853 | -83.8267
4.0856 | -83.828
4.0859 | -83.8298
4.0862 | -83.8315
4.0865 | -83.8324
4.0869 | -83.8329
4.0872 | -83.8343
4.0875 | -83.8344
4.0878 | -83.8352
4.0881 | -83.8361
4.0884 | -83.8367
4.0887 | -83.8379
4.089 | -83.8385
4.0893 | -83.8396
4.0896 | -83.841
4.0899 | -83.8428
4.0902 | -83.8452
4.0906 | -83.8469
4.0909 | -83.8483
4.0912 | -83.8502
4.0915 | -83.8511
4.0918 | -83.8519
4.0921 | -83.8529
4.0924 | -83.8539
4.0927 | -83.8553
4.093 | -83.8569
4.0933 | -83.8583
4.0936 | -83.8592
4.094 | -83.86
4.0943 | -83.8615
4.0946 | -83.8626
4.0949 | -83.8635
4.0952 | -83.8643
4.0955 | -83.8654
4.0958 | -83.8668
4.0961 | -83.8679
4.0964 | -83.8685
4.0967 | -83.8693
4.097 | -83.871
4.0973 | -83.8714
4.0977 | -83.8726
4.098 | -83.8751
4.0983 | -83.8771
4.0986 | -83.8782
4.0989 | -83.8787
4.0992 | -83.8789
4.0995 | -83.88
4.0998 | -83.8812
4.1001 | -83.8823
4.1004 | -83.8836
4.1007 | -83.8851
4.1011 | -83.8859
4.1014 | -83.8877
4.1017 | -83.8891
4.102 | -83.8909
4.1023 | -83.8919
4.1026 | -83.8934
4.1029 | -83.8948
4.1032 | -83.8963
4.1035 | -83.8978
4.1038 | -83.8988
4.1041 | -83.8997
4.1044 | -83.901
4.1048 | -83.9031
4.1051 | -83.9056
4.1054 | -83.9091
4.1057 | -83.9121
4.106 | -83.9116
4.1063 | -83.91
4.1066 | -83.9071
4.1069 | -83.9086
4.1072 | -83.9101
4.1075 | -83.9126
4.1078 | -83.9139
4.1082 | -83.9152
4.1085 | -83.9169
4.1088 | -83.9187
4.1091 | -83.9199
4.1094 | -83.9216
4.1097 | -83.9244
4.11 | -83.9256
4.1103 | -83.9273
4.1106 | -83.9287
4.1109 | -83.9294
4.1112 | -83.9296
4.1116 | -83.9309
4.1119 | -83.9324
4.1122 | -83.9339
4.1125 | -83.9365
4.1128 | -83.9383
4.1131 | -83.9399
4.1134 | -83.942
4.1137 | -83.9436
4.114 | -83.9453
4.1143 | -83.9477
4.1146 | -83.9493
4.1149 | -83.9508
4.1153 | -83.9526
4.1156 | -83.9549
4.1159 | -83.9556
4.1162 | -83.9568
4.1165 | -83.9589
4.1168 | -83.9605
4.1171 | -83.9624
4.1174 | -83.964
4.1177 | -83.9652
4.118 | -83.9669
4.1183 | -83.9676
4.1187 | -83.9688
4.119 | -83.9711
4.1193 | -83.973
4.1196 | -83.9749
4.1199 | -83.9769
4.1202 | -83.9782
4.1205 | -83.9799
4.1208 | -83.982
4.1211 | -83.9837
4.1214 | -83.985
4.1217 | -83.986
4.122 | -83.9868
4.1224 | -83.987
4.1227 | -83.9883
4.123 | -83.9897
4.1233 | -83.9905
4.1236 | -83.9914
4.1239 | -83.9936
4.1242 | -83.997
4.1245 | -83.9994
4.1248 | -84.0022
4.1251 | -84.0036
4.1254 | -84.0038
4.1258 | -84.0038
4.1261 | -84.0053
4.1264 | -84.0068
4.1267 | -84.008
4.127 | -84.0095
4.1273 | -84.0109
4.1276 | -84.0122
4.1279 | -84.0129
4.1282 | -84.0141
4.1285 | -84.0158
4.1288 | -84.0173
4.1292 | -84.0187
4.1295 | -84.0202
4.1298 | -84.0216
4.1301 | -84.0231
4.1304 | -84.0244
4.1307 | -84.0257
4.131 | -84.027
4.1313 | -84.0288
4.1316 | -84.0306
4.1319 | -84.0308
4.1322 | -84.0317
4.1325 | -84.0339
4.1329 | -84.036
4.1332 | -84.0384
4.1335 | -84.0406
4.1338 | -84.042
4.1341 | -84.0427
4.1344 | -84.0442
4.1347 | -84.0459
4.135 | -84.0472
4.1353 | -84.0485
4.1356 | -84.0502
4.1359 | -84.0518
4.1363 | -84.0534
4.1366 | -84.0551
4.1369 | -84.0568
4.1372 | -84.0585
4.1375 | -84.0604
4.1378 | -84.0616
4.1381 | -84.0631
4.1384 | -84.0636
4.1387 | -84.0629
4.139 | -84.0636
4.1393 | -84.0655
4.1396 | -84.0677
4.14 | -84.0691
4.1403 | -84.0723
4.1406 | -84.0764
4.1409 | -84.0788
4.1412 | -84.0806
4.1415 | -84.0811
4.1418 | -84.0821
4.1421 | -84.0842
4.1424 | -84.0853
4.1427 | -84.0863
4.143 | -84.0874
4.1434 | -84.089
4.1437 | -84.0903
4.144 | -84.0919
4.1443 | -84.0936
4.1446 | -84.0955
4.1449 | -84.0981
4.1452 | -84.1007
4.1455 | -84.1026
4.1458 | -84.1043
4.1461 | -84.1058
4.1464 | -84.1075
4.1468 | -84.1092
4.1471 | -84.1124
4.1474 | -84.1145
4.1477 | -84.1156
4.148 | -84.1174
4.1483 | -84.1191
4.1486 | -84.123
4.1489 | -84.1262
4.1492 | -84.1279
4.1495 | -84.1287
4.1498 | -84.1298
4.1501 | -84.1317
4.1505 | -84.1341
4.1508 | -84.1366
4.1511 | -84.1386
4.1514 | -84.141
4.1517 | -84.1424
4.152 | -84.144
4.1523 | -84.1464
4.1526 | -84.1488
4.1529 | -84.151
4.1532 | -84.1526
4.1535 | -84.1544
4.1539 | -84.1555
4.1542 | -84.1562
4.1545 | -84.1574
4.1548 | -84.1599
4.1551 | -84.1621
4.1554 | -84.1632
4.1557 | -84.1655
4.156 | -84.1672
4.1563 | -84.169
4.1566 | -84.1712
4.1569 | -84.1731
4.1572 | -84.1753
4.1576 | -84.1781
4.1579 | -84.1799
4.1582 | -84.1811
4.1585 | -84.1818
4.1588 | -84.1833
4.1591 | -84.185
4.1594 | -84.1888
4.1597 | -84.1925
4.16 | -84.1945
4.1603 | -84.1971
4.1606 | -84.1983
4.161 | -84.201
4.1613 | -84.2036
4.1616 | -84.2051
4.1619 | -84.2067
4.1622 | -84.2083
4.1625 | -84.2103
4.1628 | -84.2116
4.1631 | -84.2139
4.1634 | -84.2159
4.1637 | -84.2178
4.164 | -84.2204
4.1644 | -84.2223
4.1647 | -84.2238
4.165 | -84.2238
4.1653 | -84.2235
4.1656 | -84.2254
4.1659 | -84.2262
4.1662 | -84.2257
4.1665 | -84.2271
4.1668 | -84.2288
4.1671 | -84.2316
4.1674 | -84.2335
4.1677 | -84.2349
4.1681 | -84.2356
4.1684 | -84.237
4.1687 | -84.2385
4.169 | -84.2411
4.1693 | -84.2437
4.1696 | -84.2455
4.1699 | -84.2466
4.1702 | -84.2484
4.1705 | -84.2511
4.1708 | -84.2532
4.1711 | -84.2548
4.1715 | -84.2558
4.1718 | -84.2569
4.1721 | -84.2586
4.1724 | -84.2608
4.1727 | -84.2613
4.173 | -84.2619
4.1733 | -84.2645
4.1736 | -84.2665
4.1739 | -84.2674
4.1742 | -84.2696
4.1745 | -84.2714
4.1748 | -84.2731
4.1752 | -84.2756
4.1755 | -84.2778
4.1758 | -84.2796
4.1761 | -84.2811
4.1764 | -84.2827
4.1767 | -84.2844
4.177 | -84.2865
4.1773 | -84.2892
4.1776 | -84.2916
4.1779 | -84.2941
4.1782 | -84.2958
4.1786 | -84.2979
4.1789 | -84.3001
4.1792 | -84.3014
4.1795 | -84.303
4.1798 | -84.3047
4.1801 | -84.3064
4.1804 | -84.3087
4.1807 | -84.3104
4.181 | -84.3128
4.1813 | -84.3149
4.1816 | -84.3165
4.182 | -84.3175
4.1823 | -84.3182
4.1826 | -84.3201
4.1829 | -84.3223
4.1832 | -84.3253
4.1835 | -84.3286
4.1838 | -84.3329
4.1841 | -84.3366
4.1844 | -84.3389
4.1847 | -84.3408
4.185 | -84.3431
4.1853 | -84.3458
4.1857 | -84.3464
4.186 | -84.3479
4.1863 | -84.3508
4.1866 | -84.3522
4.1869 | -84.3534
4.1872 | -84.3548
4.1875 | -84.3564
4.1878 | -84.3583
4.1881 | -84.3604
4.1884 | -84.3622
4.1887 | -84.364
4.1891 | -84.3655
4.1894 | -84.3673
4.1897 | -84.3691
4.19 | -84.3703
4.1903 | -84.3711
4.1906 | -84.3721
4.1909 | -84.374
4.1912 | -84.3768
4.1915 | -84.3788
4.1918 | -84.3796
4.1921 | -84.3806
4.1924 | -84.3821
4.1928 | -84.3834
4.1931 | -84.385
4.1934 | -84.3866
4.1937 | -84.3885
4.194 | -84.3901
4.1943 | -84.3915
4.1946 | -84.3926
4.1949 | -84.3937
4.1952 | -84.3951
4.1955 | -84.3984
4.1958 | -84.3994
4.1962 | -84.4029
4.1965 | -84.4062
4.1968 | -84.4083
4.1971 | -84.4109
4.1974 | -84.4126
4.1977 | -84.415
4.198 | -84.4178
4.1983 | -84.4204
4.1986 | -84.4229
4.1989 | -84.4251
4.1992 | -84.4265
4.1996 | -84.427
4.1999 | -84.4293
4.2002 | -84.4325
4.2005 | -84.435
4.2008 | -84.4365
4.2011 | -84.4375
4.2014 | -84.4398
4.2017 | -84.4423
4.202 | -84.4442
4.2023 | -84.4462
4.2026 | -84.4485
4.2029 | -84.4505
4.2033 | -84.4527
4.2036 | -84.4566
4.2039 | -84.4604
4.2042 | -84.4633
4.2045 | -84.4661
4.2048 | -84.4678
4.2051 | -84.4702
4.2054 | -84.4718
4.2057 | -84.4736
4.206 | -84.4752
4.2063 | -84.4776
4.2067 | -84.4803
4.207 | -84.4839
4.2073 | -84.4868
4.2076 | -84.4893
4.2079 | -84.4918
4.2082 | -84.494
4.2085 | -84.4966
4.2088 | -84.4984
4.2091 | -84.5
4.2094 | -84.5017
4.2097 | -84.5045
4.21 | -84.5074
4.2104 | -84.5089
4.2107 | -84.5097
4.211 | -84.5109
4.2113 | -84.5117
4.2116 | -84.5136
4.2119 | -84.5166
4.2122 | -84.5195
4.2125 | -84.5213
4.2128 | -84.5232
4.2131 | -84.5245
4.2134 | -84.5252
4.2138 | -84.5273
4.2141 | -84.5281
4.2144 | -84.5288
4.2147 | -84.5309
4.215 | -84.5329
4.2153 | -84.534
4.2156 | -84.5349
4.2159 | -84.5364
4.2162 | -84.5386
4.2165 | -84.5402
4.2168 | -84.5416
4.2172 | -84.5427
4.2175 | -84.544
4.2178 | -84.5453
4.2181 | -84.5466
4.2184 | -84.5486
4.2187 | -84.551
4.219 | -84.553
4.2193 | -84.555
4.2196 | -84.5567
4.2199 | -84.5583
4.2202 | -84.5615
4.2205 | -84.5638
4.2209 | -84.5651
4.2212 | -84.5663
4.2215 | -84.5675
4.2218 | -84.5678
4.2221 | -84.5678
4.2224 | -84.5696
4.2227 | -84.5718
4.223 | -84.5734
4.2233 | -84.575
4.2236 | -84.5762
4.2239 | -84.578
4.2243 | -84.5801
4.2246 | -84.5819
4.2249 | -84.5843
4.2252 | -84.5857
4.2255 | -84.5866
4.2258 | -84.5878
4.2261 | -84.589
4.2264 | -84.5904
4.2267 | -84.5919
4.227 | -84.5932
4.2273 | -84.5945
4.2276 | -84.5961
4.228 | -84.5977
4.2283 | -84.6003
4.2286 | -84.6022
4.2289 | -84.6037
4.2292 | -84.605
4.2295 | -84.6062
4.2298 | -84.6079
4.2301 | -84.6096
4.2304 | -84.6112
4.2307 | -84.6127
4.231 | -84.6146
4.2314 | -84.6167
4.2317 | -84.6186
4.232 | -84.6213
4.2323 | -84.6244
4.2326 | -84.6264
4.2329 | -84.6278
4.2332 | -84.6289
4.2335 | -84.6301
4.2338 | -84.6317
4.2341 | -84.6331
4.2344 | -84.6353
4.2348 | -84.6368
4.2351 | -84.6382
4.2354 | -84.6413
4.2357 | -84.6442
4.236 | -84.6454
4.2363 | -84.6477
4.2366 | -84.65
4.2369 | -84.6515
4.2372 | -84.6526
4.2375 | -84.654
4.2378 | -84.6557
4.2381 | -84.6565
4.2385 | -84.6581
4.2388 | -84.66
4.2391 | -84.6618
4.2394 | -84.6635
4.2397 | -84.6647
4.24 | -84.6656
4.2403 | -84.6676
4.2406 | -84.6687
4.2409 | -84.6714
4.2412 | -84.6733
4.2415 | -84.6748
4.2419 | -84.6764
4.2422 | -84.6785
4.2425 | -84.6808
4.2428 | -84.6827
4.2431 | -84.6853
4.2434 | -84.6872
4.2437 | -84.689
4.244 | -84.6898
4.2443 | -84.6907
4.2446 | -84.6921
4.2449 | -84.6932
4.2452 | -84.6941
4.2456 | -84.695
4.2459 | -84.6964
4.2462 | -84.6977
4.2465 | -84.698
4.2468 | -84.6997
4.2471 | -84.7012
4.2474 | -84.7031
4.2477 | -84.7048
4.248 | -84.7068
4.2483 | -84.7083
4.2486 | -84.7105
4.249 | -84.7122
4.2493 | -84.7125
4.2496 | -84.7137
4.2499 | -84.7152
4.2502 | -84.7164
4.2505 | -84.718
4.2508 | -84.721
4.2511 | -84.7232
4.2514 | -84.7248
4.2517 | -84.727
4.252 | -84.7288
4.2523 | -84.7304
4.2527 | -84.7317
4.253 | -84.7336
4.2533 | -84.7369
4.2536 | -84.7379
4.2539 | -84.7394
4.2542 | -84.7411
4.2545 | -84.7432
4.2548 | -84.7475
4.2551 | -84.7522
4.2554 | -84.7543
4.2557 | -84.7556
4.2561 | -84.7573
4.2564 | -84.7585
4.2567 | -84.7587
4.257 | -84.7596
4.2573 | -84.7617
4.2576 | -84.763
4.2579 | -84.7638
4.2582 | -84.7652
4.2585 | -84.7663
4.2588 | -84.7676
4.2591 | -84.7686
4.2595 | -84.7702
4.2598 | -84.7727
4.2601 | -84.7753
4.2604 | -84.777
4.2607 | -84.7786
4.261 | -84.7801
4.2613 | -84.7813
4.2616 | -84.7822
4.2619 | -84.7825
4.2622 | -84.7831
4.2625 | -84.7844
4.2628 | -84.7858
4.2632 | -84.7876
4.2635 | -84.7891
4.2638 | -84.7899
4.2641 | -84.7907
4.2644 | -84.7919
4.2647 | -84.7929
4.265 | -84.7944
4.2653 | -84.7957
4.2656 | -84.7972
4.2659 | -84.7987
4.2662 | -84.7995
4.2666 | -84.801
4.2669 | -84.8025
4.2672 | -84.8033
4.2675 | -84.8042
4.2678 | -84.8061
4.2681 | -84.8071
4.2684 | -84.8087
4.2687 | -84.8103
4.269 | -84.8109
4.2693 | -84.8109
4.2696 | -84.8123
4.2699 | -84.8136
4.2703 | -84.815
4.2706 | -84.8161
4.2709 | -84.8174
4.2712 | -84.8183
4.2715 | -84.8193
4.2718 | -84.8209
4.2721 | -84.8229
4.2724 | -84.8239
4.2727 | -84.8246
4.273 | -84.8265
4.2733 | -84.8281
4.2737 | -84.8292
4.274 | -84.8305
4.2743 | -84.8315
4.2746 | -84.8327
4.2749 | -84.8341
4.2752 | -84.8341
4.2755 | -84.835
4.2758 | -84.836
4.2761 | -84.8363
4.2764 | -84.8381
4.2767 | -84.8392
4.2771 | -84.8411
4.2774 | -84.8421
4.2777 | -84.8442
4.278 | -84.8465
4.2783 | -84.8489
4.2786 | -84.8508
4.2789 | -84.8521
4.2792 | -84.8536
4.2795 | -84.8539
4.2798 | -84.8545
4.2801 | -84.856
4.2804 | -84.8572
4.2808 | -84.8584
4.2811 | -84.8601
4.2814 | -84.8608
4.2817 | -84.8621
4.282 | -84.8632
4.2823 | -84.8645
4.2826 | -84.8659
4.2829 | -84.8684
4.2832 | -84.8696
4.2835 | -84.871
4.2838 | -84.8715
4.2842 | -84.8731
4.2845 | -84.8748
4.2848 | -84.876
4.2851 | -84.8777
4.2854 | -84.8789
4.2857 | -84.8801
4.286 | -84.8811
4.2863 | -84.8822
4.2866 | -84.8843
4.2869 | -84.8852
4.2872 | -84.8856
4.2875 | -84.8868
4.2879 | -84.889
4.2882 | -84.892
4.2885 | -84.894
4.2888 | -84.8956
4.2891 | -84.8964
4.2894 | -84.8972
4.2897 | -84.8994
4.29 | -84.9019
4.2903 | -84.9026
4.2906 | -84.9025
4.2909 | -84.9038
4.2913 | -84.9055
4.2916 | -84.9057
4.2919 | -84.906
4.2922 | -84.9076
4.2925 | -84.9089
4.2928 | -84.9101
4.2931 | -84.9116
4.2934 | -84.913
4.2937 | -84.9112
4.294 | -84.9135
4.2943 | -84.9145
4.2947 | -84.916
4.295 | -84.9174
4.2953 | -84.9185
4.2956 | -84.9201
4.2959 | -84.9215
4.2962 | -84.922
4.2965 | -84.9237
4.2968 | -84.9261
4.2971 | -84.9286
4.2974 | -84.9295
4.2977 | -84.9313
4.298 | -84.9327
4.2984 | -84.9318
4.2987 | -84.9299
4.299 | -84.9303
4.2993 | -84.9315
4.2996 | -84.9338
4.2999 | -84.935
4.3002 | -84.9356
4.3005 | -84.9365
4.3008 | -84.9375
4.3011 | -84.9377
4.3014 | -84.9374
4.3018 | -84.9389
4.3021 | -84.9402
4.3024 | -84.9409
4.3027 | -84.941
4.303 | -84.9427
4.3033 | -84.9443
4.3036 | -84.945
4.3039 | -84.9459
4.3042 | -84.9462
4.3045 | -84.9476
4.3048 | -84.9497
4.3051 | -84.951
4.3055 | -84.9524
4.3058 | -84.9527
4.3061 | -84.953
4.3064 | -84.9539
4.3067 | -84.9557
4.307 | -84.957
4.3073 | -84.9584
4.3076 | -84.9597
4.3079 | -84.9594
4.3082 | -84.9584
4.3085 | -84.9569
4.3089 | -84.9568
4.3092 | -84.9577
4.3095 | -84.9584
4.3098 | -84.959
4.3101 | -84.9596
4.3104 | -84.9584
4.3107 | -84.9567
4.311 | -84.9575
4.3113 | -84.958
4.3116 | -84.9584
4.3119 | -84.9603
4.3123 | -84.9626
4.3126 | -84.9624
4.3129 | -84.9621
4.3132 | -84.9625
4.3135 | -84.9628
4.3138 | -84.9634
4.3141 | -84.9648
4.3144 | -84.9652
4.3147 | -84.9642
4.315 | -84.9632
4.3153 | -84.9641
4.3156 | -84.9649
4.316 | -84.9655
4.3163 | -84.9658
4.3166 | -84.9665
4.3169 | -84.9668
4.3172 | -84.9689
4.3175 | -84.9699
4.3178 | -84.9696
4.3181 | -84.9706
4.3184 | -84.9722
4.3187 | -84.9727
4.319 | -84.9722
4.3194 | -84.9728
4.3197 | -84.9727
4.32 | -84.9727
4.3203 | -84.9737
4.3206 | -84.9744
4.3209 | -84.9756
4.3212 | -84.9757
4.3215 | -84.9759
4.3218 | -84.9759
4.3221 | -84.9758
4.3224 | -84.9761
4.3227 | -84.9764
4.3231 | -84.9767
4.3234 | -84.9761
4.3237 | -84.9758
4.324 | -84.9772
4.3243 | -84.9795
4.3246 | -84.9803
4.3249 | -84.9815
4.3252 | -84.9819
4.3255 | -84.9824
4.3258 | -84.9826
4.3261 | -84.9831
4.3265 | -84.9838
4.3268 | -84.9835
4.3271 | -84.9835
4.3274 | -84.9835
4.3277 | -84.9839
4.328 | -84.9837
4.3283 | -84.9836
4.3286 | -84.9832
4.3289 | -84.9827
4.3292 | -84.9834
4.3295 | -84.9831
4.3299 | -84.9841
4.3302 | -84.9849
4.3305 | -84.9849
4.3308 | -84.985
4.3311 | -84.9853
4.3314 | -84.9856
4.3317 | -84.9879
4.332 | -84.988
4.3323 | -84.9876
4.3326 | -84.9865
4.3329 | -84.9861
4.3332 | -84.987
4.3336 | -84.9879
4.3339 | -84.9887
4.3342 | -84.9891
4.3345 | -84.9888
4.3348 | -84.9891
4.3351 | -84.989
4.3354 | -84.9886
4.3357 | -84.9898
4.336 | -84.99
4.3363 | -84.9892
4.3366 | -84.989
4.337 | -84.9889
4.3373 | -84.9881
4.3376 | -84.9861
4.3379 | -84.9852
4.3382 | -84.9856
4.3385 | -84.9853
4.3388 | -84.9859
4.3391 | -84.9859
4.3394 | -84.9858
4.3397 | -84.9862
4.34 | -84.9864
4.3403 | -84.9863
4.3407 | -84.9866
4.341 | -84.9875
4.3413 | -84.988
4.3416 | -84.9878
4.3419 | -84.9875
4.3422 | -84.9878
4.3425 | -84.9873
4.3428 | -84.9842
4.3431 | -84.984
4.3434 | -84.9866
4.3437 | -84.9884
4.3441 | -84.988
4.3444 | -84.9879
4.3447 | -84.9876
4.345 | -84.9871
4.3453 | -84.987
4.3456 | -84.9871
4.3459 | -84.9857
4.3462 | -84.9842
4.3465 | -84.9849
4.3468 | -84.985
4.3471 | -84.9839
4.3475 | -84.9835
4.3478 | -84.983
4.3481 | -84.9825
4.3484 | -84.9817
4.3487 | -84.9812
4.349 | -84.981
4.3493 | -84.9808
4.3496 | -84.9802
4.3499 | -84.9792
4.3502 | -84.9786
4.3505 | -84.9784
4.3508 | -84.9769
4.3512 | -84.976
4.3515 | -84.975
4.3518 | -84.9747
4.3521 | -84.9742
4.3524 | -84.9723
4.3527 | -84.9714
4.353 | -84.9713
4.3533 | -84.9714
4.3536 | -84.9715
4.3539 | -84.9715
4.3542 | -84.9707
4.3546 | -84.969
4.3549 | -84.9684
4.3552 | -84.9683
4.3555 | -84.9684
4.3558 | -84.9693
4.3561 | -84.9671
4.3564 | -84.9676
4.3567 | -84.9668
4.357 | -84.9662
4.3573 | -84.965
4.3576 | -84.9635
4.3579 | -84.9627
4.3583 | -84.9622
4.3586 | -84.9618
4.3589 | -84.9612
4.3592 | -84.9602
4.3595 | -84.9586
4.3598 | -84.9565
4.3601 | -84.9562
4.3604 | -84.9556
4.3607 | -84.954
4.361 | -84.9521
4.3613 | -84.9502
4.3617 | -84.9495
4.362 | -84.9488
4.3623 | -84.9478
4.3626 | -84.9459
4.3629 | -84.9434
4.3632 | -84.943
4.3635 | -84.9431
4.3638 | -84.9413
4.3641 | -84.9394
4.3644 | -84.9381
4.3647 | -84.9369
4.3651 | -84.9358
4.3654 | -84.9332
4.3657 | -84.9312
4.366 | -84.9317
4.3663 | -84.9319
4.3666 | -84.9313
4.3669 | -84.9302
4.3672 | -84.9287
4.3675 | -84.9277
4.3678 | -84.9264
4.3681 | -84.9254
4.3684 | -84.9244
4.3688 | -84.9239
4.3691 | -84.923
4.3694 | -84.9219
4.3697 | -84.9209
4.37 | -84.9221
4.3703 | -84.9215
4.3706 | -84.9206
4.3709 | -84.9197
4.3712 | -84.9186
4.3715 | -84.917
4.3718 | -84.9147
4.3722 | -84.9133
4.3725 | -84.9122
4.3728 | -84.9113
4.3731 | -84.9111
4.3734 | -84.9104
4.3737 | -84.9094
4.374 | -84.9072
4.3743 | -84.9064
4.3746 | -84.9055
4.3749 | -84.9048
4.3752 | -84.9027
4.3755 | -84.8982
4.3759 | -84.8958
4.3762 | -84.8946
4.3765 | -84.8927
4.3768 | -84.8898
4.3771 | -84.8884
4.3774 | -84.8865
4.3777 | -84.8846
4.378 | -84.8825
4.3783 | -84.8798
4.3786 | -84.8783
4.3789 | -84.8773
4.3793 | -84.8759
4.3796 | -84.8743
4.3799 | -84.8719
4.3802 | -84.8701
4.3805 | -84.8689
4.3808 | -84.8681
4.3811 | -84.8667
4.3814 | -84.8669
4.3817 | -84.8657
4.382 | -84.8636
4.3823 | -84.8609
4.3826 | -84.8595
4.383 | -84.8578
4.3833 | -84.857
4.3836 | -84.8569
4.3839 | -84.8564
4.3842 | -84.8543
4.3845 | -84.8515
4.3848 | -84.8501
4.3851 | -84.8482
4.3854 | -84.8464
4.3857 | -84.845
4.386 | -84.8435
4.3864 | -84.842
4.3867 | -84.8398
4.387 | -84.8375
4.3873 | -84.8356
4.3876 | -84.8338
4.3879 | -84.8315
4.3882 | -84.8297
4.3885 | -84.8286
4.3888 | -84.8273
4.3891 | -84.8254
4.3894 | -84.823
4.3898 | -84.8201
4.3901 | -84.8163
4.3904 | -84.811
4.3907 | -84.8071
4.391 | -84.806
4.3913 | -84.805
4.3916 | -84.8032
4.3919 | -84.8017
4.3922 | -84.8006
4.3925 | -84.7985
4.3928 | -84.7962
4.3931 | -84.7936
4.3935 | -84.7909
4.3938 | -84.7885
4.3941 | -84.7873
4.3944 | -84.7868
4.3947 | -84.7855
4.395 | -84.7842
4.3953 | -84.783
4.3956 | -84.7814
4.3959 | -84.7796
4.3962 | -84.7779
4.3965 | -84.7766
4.3969 | -84.7749
4.3972 | -84.7726
4.3975 | -84.7713
4.3978 | -84.7702
4.3981 | -84.7675
4.3984 | -84.7649
4.3987 | -84.7633
4.399 | -84.762
4.3993 | -84.7605
4.3996 | -84.7587
4.3999 | -84.7576
4.4002 | -84.7561
4.4006 | -84.7549
4.4009 | -84.7523
4.4012 | -84.7506
4.4015 | -84.7512
4.4018 | -84.7527
4.4021 | -84.7537
4.4024 | -84.7518
4.4027 | -84.7495
4.403 | -84.7476
4.4033 | -84.7457
4.4036 | -84.7438
4.404 | -84.7425
4.4043 | -84.7408
4.4046 | -84.7385
4.4049 | -84.7367
4.4052 | -84.7349
4.4055 | -84.7329
4.4058 | -84.7305
4.4061 | -84.728
4.4064 | -84.726
4.4067 | -84.7241
4.407 | -84.7234
4.4074 | -84.7217
4.4077 | -84.7186
4.408 | -84.7153
4.4083 | -84.7123
4.4086 | -84.7107
4.4089 | -84.7094
4.4092 | -84.7074
4.4095 | -84.7059
4.4098 | -84.703
4.4101 | -84.7006
4.4104 | -84.6987
4.4107 | -84.6966
4.4111 | -84.6944
4.4114 | -84.6922
4.4117 | -84.6903
4.412 | -84.6887
4.4123 | -84.6874
4.4126 | -84.6851
4.4129 | -84.6829
4.4132 | -84.6815
4.4135 | -84.6797
4.4138 | -84.6778
4.4141 | -84.6759
4.4145 | -84.6735
4.4148 | -84.6725
4.4151 | -84.6708
4.4154 | -84.6689
4.4157 | -84.6672
4.416 | -84.6653
4.4163 | -84.6638
4.4166 | -84.6621
4.4169 | -84.6603
4.4172 | -84.6587
4.4175 | -84.6571
4.4178 | -84.6553
4.4182 | -84.6537
4.4185 | -84.6517
4.4188 | -84.6507
4.4191 | -84.65
4.4194 | -84.6483
4.4197 | -84.6464
4.42 | -84.6451
4.4203 | -84.6443
4.4206 | -84.6423
4.4209 | -84.6405
4.4212 | -84.6387
4.4216 | -84.6371
4.4219 | -84.6355
4.4222 | -84.6341
4.4225 | -84.6334
4.4228 | -84.6324
4.4231 | -84.6307
4.4234 | -84.6294
4.4237 | -84.6282
4.424 | -84.6264
4.4243 | -84.6245
4.4246 | -84.6234
4.425 | -84.6223
4.4253 | -84.621
4.4256 | -84.6194
4.4259 | -84.618
4.4262 | -84.6159
4.4265 | -84.6134
4.4268 | -84.6115
4.4271 | -84.6097
4.4274 | -84.6094
4.4277 | -84.6081
4.428 | -84.6065
4.4283 | -84.6045
4.4287 | -84.6024
4.429 | -84.601
4.4293 | -84.6006
4.4296 | -84.5997
4.4299 | -84.5979
4.4302 | -84.5964
4.4305 | -84.5949
4.4308 | -84.5933
4.4311 | -84.5916
4.4314 | -84.5903
4.4317 | -84.5895
4.4321 | -84.5882
4.4324 | -84.5861
4.4327 | -84.5843
4.433 | -84.5827
4.4333 | -84.5812
4.4336 | -84.5779
4.4339 | -84.5747
4.4342 | -84.5724
4.4345 | -84.5706
4.4348 | -84.5688
4.4351 | -84.5672
4.4354 | -84.5662
4.4358 | -84.5651
4.4361 | -84.5651
4.4364 | -84.5635
4.4367 | -84.5642
4.437 | -84.5638
4.4373 | -84.5624
4.4376 | -84.5609
4.4379 | -84.56
4.4382 | -84.5595
4.4385 | -84.5586
4.4388 | -84.5579
4.4392 | -84.5583
4.4395 | -84.5593
4.4398 | -84.5585
4.4401 | -84.5574
4.4404 | -84.5569
4.4407 | -84.5557
4.441 | -84.5542
4.4413 | -84.5525
4.4416 | -84.551
4.4419 | -84.5504
4.4422 | -84.5495
4.4426 | -84.5482
4.4429 | -84.5467
4.4432 | -84.5453
4.4435 | -84.5439
4.4438 | -84.5425
4.4441 | -84.5412
4.4444 | -84.5402
4.4447 | -84.5377
4.445 | -84.536
4.4453 | -84.5364
4.4456 | -84.5358
4.4459 | -84.5357
4.4463 | -84.5356
4.4466 | -84.5346
4.4469 | -84.5329
4.4472 | -84.5309
4.4475 | -84.5299
4.4478 | -84.5286
4.4481 | -84.5287
4.4484 | -84.5282
4.4487 | -84.5274
4.449 | -84.5264
4.4493 | -84.5256
4.4497 | -84.5229
4.45 | -84.5194
4.4503 | -84.5177
4.4506 | -84.5169
4.4509 | -84.515
4.4512 | -84.5136
4.4515 | -84.513
4.4518 | -84.5119
4.4521 | -84.5104
4.4524 | -84.5079
4.4527 | -84.5065
4.453 | -84.5064
4.4534 | -84.5058
4.4537 | -84.5048
4.454 | -84.5047
4.4543 | -84.5052
4.4546 | -84.5048
4.4549 | -84.5038
4.4552 | -84.5031
4.4555 | -84.5026
4.4558 | -84.5022
4.4561 | -84.502
4.4564 | -84.5014
4.4568 | -84.5007
4.4571 | -84.5007
4.4574 | -84.5007
4.4577 | -84.5009
4.458 | -84.5006
4.4583 | -84.4999
4.4586 | -84.4995
4.4589 | -84.4985
4.4592 | -84.4978
4.4595 | -84.497
4.4598 | -84.496
4.4601 | -84.4959
4.4605 | -84.4954
4.4608 | -84.4948
4.4611 | -84.4946
4.4614 | -84.4943
4.4617 | -84.4941
4.462 | -84.4933
4.4623 | -84.4931
4.4626 | -84.4915
4.4629 | -84.4905
4.4632 | -84.4909
4.4635 | -84.4906
4.4639 | -84.4901
4.4642 | -84.4895
4.4645 | -84.4892
4.4648 | -84.49
4.4651 | -84.4911
4.4654 | -84.4907
4.4657 | -84.4912
4.466 | -84.493
4.4663 | -84.4929
4.4666 | -84.4931
4.4669 | -84.4931
4.4673 | -84.4927
4.4676 | -84.493
4.4679 | -84.4935
4.4682 | -84.4938
4.4685 | -84.4938
4.4688 | -84.4936
4.4691 | -84.4934
4.4694 | -84.4932
4.4697 | -84.4938
4.47 | -84.4943
4.4703 | -84.4941
4.4706 | -84.4937
4.471 | -84.4936
4.4713 | -84.4936
4.4716 | -84.4943
4.4719 | -84.4953
4.4722 | -84.4953
4.4725 | -84.4951
4.4728 | -84.4955
4.4731 | -84.4973
4.4734 | -84.4985
4.4737 | -84.4981
4.474 | -84.4991
4.4744 | -84.5008
4.4747 | -84.501
4.475 | -84.5009
4.4753 | -84.501
4.4756 | -84.5009
4.4759 | -84.5008
4.4762 | -84.5003
4.4765 | -84.4996
4.4768 | -84.4995
4.4771 | -84.5005
4.4774 | -84.5022
4.4778 | -84.5035
4.4781 | -84.5033
4.4784 | -84.5032
4.4787 | -84.5034
4.479 | -84.5036
4.4793 | -84.5038
4.4796 | -84.5038
4.4799 | -84.5039
4.4802 | -84.5043
4.4805 | -84.5045
4.4808 | -84.506
4.4811 | -84.5061
4.4815 | -84.5063
4.4818 | -84.5067
4.4821 | -84.5069
4.4824 | -84.5063
4.4827 | -84.507
4.483 | -84.5083
4.4833 | -84.5102
4.4836 | -84.5111
4.4839 | -84.5117
4.4842 | -84.5124
4.4845 | -84.5123
4.4849 | -84.5117
4.4852 | -84.512
4.4855 | -84.5125
4.4858 | -84.5125
4.4861 | -84.5127
4.4864 | -84.5137
4.4867 | -84.5157
4.487 | -84.5172
4.4873 | -84.5171
4.4876 | -84.5168
4.4879 | -84.5175
4.4882 | -84.5185
4.4886 | -84.5192
4.4889 | -84.5192
4.4892 | -84.5189
4.4895 | -84.5192
4.4898 | -84.5212
4.4901 | -84.5217
4.4904 | -84.522
4.4907 | -84.5217
4.491 | -84.5214
4.4913 | -84.5221
4.4916 | -84.5233
4.492 | -84.5244
4.4923 | -84.5257
4.4926 | -84.5269
4.4929 | -84.5278
4.4932 | -84.5289
4.4935 | -84.5303
4.4938 | -84.5311
4.4941 | -84.5317
4.4944 | -84.5324
4.4947 | -84.5328
4.495 | -84.5325
4.4954 | -84.5316
4.4957 | -84.5315
4.496 | -84.5325
4.4963 | -84.5341
4.4966 | -84.5351
4.4969 | -84.536
4.4972 | -84.5369
4.4975 | -84.5379
4.4978 | -84.5389
4.4981 | -84.5403
4.4984 | -84.5413
4.4987 | -84.5428
4.4991 | -84.5435
4.4994 | -84.5442
4.4997 | -84.545
4.5 | -84.5454
4.5003 | -84.5463
4.5006 | -84.547
4.5009 | -84.5476
4.5012 | -84.5479
4.5015 | -84.5488
4.5018 | -84.5505
4.5021 | -84.5508
4.5025 | -84.5516
4.5028 | -84.5526
4.5031 | -84.5541
4.5034 | -84.5559
4.5037 | -84.5572
4.504 | -84.5582
4.5043 | -84.559
4.5046 | -84.5598
4.5049 | -84.5606
4.5052 | -84.561
4.5055 | -84.5621
4.5058 | -84.5631
4.5062 | -84.5634
4.5065 | -84.5643
4.5068 | -84.5653
4.5071 | -84.5659
4.5074 | -84.5668
4.5077 | -84.5694
4.508 | -84.5708
4.5083 | -84.5721
4.5086 | -84.5733
4.5089 | -84.574
4.5092 | -84.5738
4.5096 | -84.5743
4.5099 | -84.575
4.5102 | -84.5762
4.5105 | -84.5779
4.5108 | -84.5794
4.5111 | -84.5808
4.5114 | -84.5824
4.5117 | -84.5839
4.512 | -84.585
4.5123 | -84.5863
4.5126 | -84.5863
4.513 | -84.5871
4.5133 | -84.588
4.5136 | -84.5898
4.5139 | -84.5916
4.5142 | -84.5934
4.5145 | -84.5936
4.5148 | -84.5941
4.5151 | -84.5953
4.5154 | -84.5965
4.5157 | -84.5975
4.516 | -84.5984
4.5163 | -84.6006
4.5167 | -84.6031
4.517 | -84.6045
4.5173 | -84.6059
4.5176 | -84.6079
4.5179 | -84.6099
4.5182 | -84.6115
4.5185 | -84.6131
4.5188 | -84.6153
4.5191 | -84.6173
4.5194 | -84.6191
4.5197 | -84.6204
4.5201 | -84.6217
4.5204 | -84.6234
4.5207 | -84.6252
4.521 | -84.6272
4.5213 | -84.6295
4.5216 | -84.6313
4.5219 | -84.6312
4.5222 | -84.6321
4.5225 | -84.6331
4.5228 | -84.6341
4.5231 | -84.6354
4.5234 | -84.637
4.5238 | -84.6386
4.5241 | -84.6402
4.5244 | -84.6419
4.5247 | -84.644
4.525 | -84.6462
4.5253 | -84.6491
4.5256 | -84.6502
4.5259 | -84.6522
4.5262 | -84.6537
4.5265 | -84.6549
4.5268 | -84.6567
4.5272 | -84.6585
4.5275 | -84.6602
4.5278 | -84.6609
4.5281 | -84.6623
4.5284 | -84.6642
4.5287 | -84.667
4.529 | -84.6689
4.5293 | -84.6711
4.5296 | -84.6728
4.5299 | -84.674
4.5302 | -84.676
4.5305 | -84.6774
4.5309 | -84.6784
4.5312 | -84.6796
4.5315 | -84.6817
4.5318 | -84.6834
4.5321 | -84.6846
4.5324 | -84.6847
4.5327 | -84.6855
4.533 | -84.6872
4.5333 | -84.6891
4.5336 | -84.6913
4.5339 | -84.6927
4.5343 | -84.6946
4.5346 | -84.6951
4.5349 | -84.6947
4.5352 | -84.6967
4.5355 | -84.698
4.5358 | -84.6999
4.5361 | -84.7017
4.5364 | -84.7039
4.5367 | -84.7058
4.537 | -84.7073
4.5373 | -84.7092
4.5377 | -84.7116
4.538 | -84.7144
4.5383 | -84.7163
4.5386 | -84.7181
4.5389 | -84.7203
4.5392 | -84.722
4.5395 | -84.7239
4.5398 | -84.7266
4.5401 | -84.7287
4.5404 | -84.7314
4.5407 | -84.7336
4.541 | -84.7352
4.5414 | -84.7374
4.5417 | -84.7393
4.542 | -84.7412
4.5423 | -84.7432
4.5426 | -84.745
4.5429 | -84.7468
4.5432 | -84.7488
4.5435 | -84.7508
4.5438 | -84.7528
4.5441 | -84.7547
4.5444 | -84.7565
4.5448 | -84.7581
4.5451 | -84.7603
4.5454 | -84.7625
4.5457 | -84.7646
4.546 | -84.7673
4.5463 | -84.769
4.5466 | -84.771
4.5469 | -84.7727
4.5472 | -84.7751
4.5475 | -84.777
4.5478 | -84.779
4.5481 | -84.781
4.5485 | -84.7827
4.5488 | -84.7848
4.5491 | -84.7867
4.5494 | -84.789
4.5497 | -84.7917
4.55 | -84.794
4.5503 | -84.7962
4.5506 | -84.7992
4.5509 | -84.8013
4.5512 | -84.803
4.5515 | -84.8048
4.5519 | -84.8075
4.5522 | -84.8096
4.5525 | -84.8112
4.5528 | -84.8134
4.5531 | -84.8145
4.5534 | -84.8155
4.5537 | -84.8173
4.554 | -84.8198
4.5543 | -84.8219
4.5546 | -84.8239
4.5549 | -84.8268
4.5553 | -84.8297
4.5556 | -84.8319
4.5559 | -84.8335
4.5562 | -84.8357
4.5565 | -84.8384
4.5568 | -84.8399
4.5571 | -84.8408
4.5574 | -84.8408
4.5577 | -84.8408
4.558 | -84.8429
4.5583 | -84.8444
4.5586 | -84.8466
4.559 | -84.8495
4.5593 | -84.8517
4.5596 | -84.8553
4.5599 | -84.8575
4.5602 | -84.8595
4.5605 | -84.8612
4.5608 | -84.8631
4.5611 | -84.8655
4.5614 | -84.8685
4.5617 | -84.8702
4.562 | -84.872
4.5624 | -84.8737
4.5627 | -84.8758
4.563 | -84.879
4.5633 | -84.8823
4.5636 | -84.8843
4.5639 | -84.8859
4.5642 | -84.8883
4.5645 | -84.8902
4.5648 | -84.8926
4.5651 | -84.8948
4.5654 | -84.8967
4.5657 | -84.899
4.5661 | -84.9011
4.5664 | -84.9031
4.5667 | -84.9048
4.567 | -84.9075
4.5673 | -84.9096
4.5676 | -84.9109
4.5679 | -84.9126
4.5682 | -84.9151
4.5685 | -84.9184
4.5688 | -84.9213
4.5691 | -84.923
4.5695 | -84.925
4.5698 | -84.9273
4.5701 | -84.9298
4.5704 | -84.9318
4.5707 | -84.9341
4.571 | -84.9366
4.5713 | -84.9381
4.5716 | -84.9401
4.5719 | -84.9418
4.5722 | -84.9438
4.5725 | -84.9459
4.5729 | -84.948
4.5732 | -84.9507
4.5735 | -84.9526
4.5738 | -84.9543
4.5741 | -84.9567
4.5744 | -84.9593
4.5747 | -84.9617
4.575 | -84.9643
4.5753 | -84.9656
4.5756 | -84.9661
4.5759 | -84.9679
4.5762 | -84.9702
4.5766 | -84.9737
4.5769 | -84.9756
4.5772 | -84.9775
4.5775 | -84.979
4.5778 | -84.9805
4.5781 | -84.9823
4.5784 | -84.9837
4.5787 | -84.9853
4.579 | -84.9873
4.5793 | -84.9891
4.5796 | -84.9911
4.58 | -84.9944
4.5803 | -84.9981
4.5806 | -85.0003
4.5809 | -85.0021
4.5812 | -85.0042
4.5815 | -85.0063
4.5818 | -85.0079
4.5821 | -85.0095
4.5824 | -85.0116
4.5827 | -85.0146
4.583 | -85.0171
4.5833 | -85.0193
4.5837 | -85.022
4.584 | -85.0242
4.5843 | -85.0267
4.5846 | -85.0287
4.5849 | -85.0304
4.5852 | -85.0325
4.5855 | -85.0347
4.5858 | -85.0366
4.5861 | -85.038
4.5864 | -85.0394
4.5867 | -85.0411
4.5871 | -85.0428
4.5874 | -85.0453
4.5877 | -85.0471
4.588 | -85.0465
4.5883 | -85.0478
4.5886 | -85.0494
4.5889 | -85.0513
4.5892 | -85.0531
4.5895 | -85.0551
4.5898 | -85.0567
4.5901 | -85.0579
4.5904 | -85.0598
4.5908 | -85.0615
4.5911 | -85.0622
4.5914 | -85.0629
4.5917 | -85.064
4.592 | -85.0669
4.5923 | -85.069
4.5926 | -85.0705
4.5929 | -85.0725
4.5932 | -85.0732
4.5935 | -85.0738
4.5938 | -85.0717
4.5942 | -85.0698
4.5945 | -85.0717
4.5948 | -85.0734
4.5951 | -85.0745
4.5954 | -85.0761
4.5957 | -85.0773
4.596 | -85.0782
4.5963 | -85.0784
4.5966 | -85.0758
4.5969 | -85.0766
4.5972 | -85.0778
4.5976 | -85.0789
4.5979 | -85.0804
4.5982 | -85.0827
4.5985 | -85.0844
4.5988 | -85.0858
4.5991 | -85.0875
4.5994 | -85.0898
4.5997 | -85.0913
4.6 | -85.0926
4.6003 | -85.0952
4.6006 | -85.0979
4.6009 | -85.1001
4.6013 | -85.1017
4.6016 | -85.1019
4.6019 | -85.103
4.6022 | -85.1043
4.6025 | -85.1068
4.6028 | -85.1093
4.6031 | -85.1101
4.6034 | -85.1119
4.6037 | -85.1132
4.604 | -85.1135
4.6043 | -85.1137
4.6047 | -85.1149
4.605 | -85.116
4.6053 | -85.117
4.6056 | -85.1174
4.6059 | -85.1186
4.6062 | -85.1199
4.6065 | -85.1212
4.6068 | -85.1224
4.6071 | -85.1243
4.6074 | -85.1268
4.6077 | -85.1316
4.6081 | -85.1368
4.6084 | -85.1389
4.6087 | -85.14
4.609 | -85.1409
4.6093 | -85.142
4.6096 | -85.1436
4.6099 | -85.1448
4.6102 | -85.1464
4.6105 | -85.1484
4.6108 | -85.1507
4.6111 | -85.1517
4.6114 | -85.1523
4.6118 | -85.1532
4.6121 | -85.1554
4.6124 | -85.1568
4.6127 | -85.1578
4.613 | -85.1588
4.6133 | -85.1601
4.6136 | -85.1615
4.6139 | -85.1631
4.6142 | -85.1641
4.6145 | -85.164
4.6148 | -85.1655
4.6152 | -85.1662
4.6155 | -85.1672
4.6158 | -85.1696
4.6161 | -85.1711
4.6164 | -85.1705
4.6167 | -85.1705
4.617 | -85.1714
4.6173 | -85.1724
4.6176 | -85.1739
4.6179 | -85.1772
4.6182 | -85.1804
4.6185 | -85.1836
4.6189 | -85.1853
4.6192 | -85.1866
4.6195 | -85.1883
4.6198 | -85.1893
4.6201 | -85.1897
4.6204 | -85.1913
4.6207 | -85.1934
4.621 | -85.1941
4.6213 | -85.1957
4.6216 | -85.1964
4.6219 | -85.197
4.6223 | -85.198
4.6226 | -85.1986
4.6229 | -85.1995
4.6232 | -85.2006
4.6235 | -85.2015
4.6238 | -85.2024
4.6241 | -85.2023
4.6244 | -85.2025
4.6247 | -85.203
4.625 | -85.2042
4.6253 | -85.2061
4.6257 | -85.2083
4.626 | -85.2089
4.6263 | -85.2095
4.6266 | -85.2103
4.6269 | -85.2111
4.6272 | -85.2123
4.6275 | -85.2145
4.6278 | -85.2152
4.6281 | -85.2144
4.6284 | -85.2138
4.6287 | -85.2155
4.629 | -85.2174
4.6294 | -85.2183
4.6297 | -85.2194
4.63 | -85.2205
4.6303 | -85.222
4.6306 | -85.2229
4.6309 | -85.2234
4.6312 | -85.2241
4.6315 | -85.2243
4.6318 | -85.2237
4.6321 | -85.2238
4.6324 | -85.2241
4.6328 | -85.225
4.6331 | -85.2256
4.6334 | -85.2262
4.6337 | -85.2266
4.634 | -85.2274
4.6343 | -85.2275
4.6346 | -85.2278
4.6349 | -85.2282
4.6352 | -85.2286
4.6355 | -85.2291
4.6358 | -85.2295
4.6361 | -85.2292
4.6365 | -85.2287
4.6368 | -85.2289
4.6371 | -85.2291
4.6374 | -85.2289
4.6377 | -85.2291
4.638 | -85.2301
4.6383 | -85.2305
4.6386 | -85.2307
4.6389 | -85.2305
4.6392 | -85.2316
4.6395 | -85.2315
4.6399 | -85.2329
4.6402 | -85.2359
4.6405 | -85.2368
4.6408 | -85.2364
4.6411 | -85.2363
4.6414 | -85.2366
4.6417 | -85.2365
4.642 | -85.2356
4.6423 | -85.2351
4.6426 | -85.2352
4.6429 | -85.2348
4.6433 | -85.2357
4.6436 | -85.2369
4.6439 | -85.238
4.6442 | -85.2386
4.6445 | -85.2398
4.6448 | -85.2404
4.6451 | -85.2407
4.6454 | -85.2409
4.6457 | -85.2409
4.646 | -85.2411
4.6463 | -85.2423
4.6466 | -85.2429
4.647 | -85.2427
4.6473 | -85.2436
4.6476 | -85.244
4.6479 | -85.2439
4.6482 | -85.2433
4.6485 | -85.2436
4.6488 | -85.2436
4.6491 | -85.2436
4.6494 | -85.2455
4.6497 | -85.2449
4.65 | -85.2441
4.6504 | -85.2432
4.6507 | -85.2431
4.651 | -85.2435
4.6513 | -85.2433
4.6516 | -85.2431
4.6519 | -85.2433
4.6522 | -85.2438
4.6525 | -85.2441
4.6528 | -85.2438
4.6531 | -85.2438
4.6534 | -85.2439
4.6537 | -85.2442
4.6541 | -85.2447
4.6544 | -85.2444
4.6547 | -85.2442
4.655 | -85.2454
4.6553 | -85.2473
4.6556 | -85.2464
4.6559 | -85.2462
4.6562 | -85.2458
4.6565 | -85.2465
4.6568 | -85.2466
4.6571 | -85.2466
4.6575 | -85.2464
4.6578 | -85.2466
4.6581 | -85.2474
4.6584 | -85.2485
4.6587 | -85.25
4.659 | -85.2479
4.6593 | -85.2478
4.6596 | -85.2481
4.6599 | -85.2486
4.6602 | -85.2475
4.6605 | -85.247
4.6608 | -85.247
4.6612 | -85.2472
4.6615 | -85.2477
4.6618 | -85.2478
4.6621 | -85.2475
4.6624 | -85.247
4.6627 | -85.247
4.663 | -85.2483
4.6633 | -85.2482
4.6636 | -85.2476
4.6639 | -85.2477
4.6642 | -85.2469
4.6646 | -85.246
4.6649 | -85.2452
4.6652 | -85.2449
4.6655 | -85.2461
4.6658 | -85.2462
4.6661 | -85.2451
4.6664 | -85.2412
4.6667 | -85.2403
4.667 | -85.2402
4.6673 | -85.2403
4.6676 | -85.2406
4.668 | -85.2405
4.6683 | -85.2403
4.6686 | -85.2399
4.6689 | -85.2397
4.6692 | -85.2392
4.6695 | -85.2394
4.6698 | -85.24
4.6701 | -85.2392
4.6704 | -85.2386
4.6707 | -85.2386
4.671 | -85.2393
4.6713 | -85.2394
4.6717 | -85.2383
4.672 | -85.238
4.6723 | -85.238
4.6726 | -85.2376
4.6729 | -85.2373
4.6732 | -85.2365
4.6735 | -85.2342
4.6738 | -85.2318
4.6741 | -85.231
4.6744 | -85.2319
4.6747 | -85.2327
4.6751 | -85.2327
4.6754 | -85.2318
4.6757 | -85.2313
4.676 | -85.231
4.6763 | -85.2311
4.6766 | -85.2318
4.6769 | -85.232
4.6772 | -85.2315
4.6775 | -85.2309
4.6778 | -85.2316
4.6781 | -85.2317
4.6784 | -85.2315
4.6788 | -85.2309
4.6791 | -85.2307
4.6794 | -85.23
4.6797 | -85.2308
4.68 | -85.2312
4.6803 | -85.2307
4.6806 | -85.2302
4.6809 | -85.2295
4.6812 | -85.2288
4.6815 | -85.2274
4.6818 | -85.2255
4.6822 | -85.2241
4.6825 | -85.2228
4.6828 | -85.2219
4.6831 | -85.2212
4.6834 | -85.2208
4.6837 | -85.2204
4.684 | -85.2196
4.6843 | -85.2189
4.6846 | -85.2182
4.6849 | -85.2176
4.6852 | -85.2162
4.6856 | -85.2149
4.6859 | -85.2146
4.6862 | -85.2148
4.6865 | -85.215
4.6868 | -85.2149
4.6871 | -85.2139
4.6874 | -85.2132
4.6877 | -85.213
4.688 | -85.2129
4.6883 | -85.2125
4.6886 | -85.2123
4.6889 | -85.2116
4.6893 | -85.2106
4.6896 | -85.2092
4.6899 | -85.2086
4.6902 | -85.2088
4.6905 | -85.2112
4.6908 | -85.2115
4.6911 | -85.2109
4.6914 | -85.2106
4.6917 | -85.2117
4.692 | -85.2112
4.6923 | -85.2113
4.6927 | -85.2111
4.693 | -85.2112
4.6933 | -85.212
4.6936 | -85.2117
4.6939 | -85.2111
4.6942 | -85.2104
4.6945 | -85.2103
4.6948 | -85.2094
4.6951 | -85.2097
4.6954 | -85.2092
4.6957 | -85.2075
4.696 | -85.2069
4.6964 | -85.2057
4.6967 | -85.204
4.697 | -85.2033
4.6973 | -85.2034
4.6976 | -85.2043
4.6979 | -85.2032
4.6982 | -85.2025
4.6985 | -85.2023
4.6988 | -85.2015
4.6991 | -85.2004
4.6994 | -85.1997
4.6998 | -85.1996
4.7001 | -85.2004
4.7004 | -85.1983
4.7007 | -85.198
4.701 | -85.1977
4.7013 | -85.1972
4.7016 | -85.1969
4.7019 | -85.1959
4.7022 | -85.1954
4.7025 | -85.1949
4.7028 | -85.1938
4.7032 | -85.1922
4.7035 | -85.1913
4.7038 | -85.1911
4.7041 | -85.1907
4.7044 | -85.1898
4.7047 | -85.189
4.705 | -85.1889
4.7053 | -85.1889
4.7056 | -85.1884
4.7059 | -85.19
4.7062 | -85.1928
4.7065 | -85.1943
4.7069 | -85.1928
4.7072 | -85.1911
4.7075 | -85.1905
4.7078 | -85.1889
4.7081 | -85.1881
4.7084 | -85.1879
4.7087 | -85.1879
4.709 | -85.1877
4.7093 | -85.1878
4.7096 | -85.1884
4.7099 | -85.1889
4.7103 | -85.1886
4.7106 | -85.1874
4.7109 | -85.1862
4.7112 | -85.1851
4.7115 | -85.1842
4.7118 | -85.1835
4.7121 | -85.1828
4.7124 | -85.1828
4.7127 | -85.1826
4.713 | -85.182
4.7133 | -85.1815
4.7136 | -85.181
4.714 | -85.1799
4.7143 | -85.1788
4.7146 | -85.1792
4.7149 | -85.1788
4.7152 | -85.178
4.7155 | -85.1773
4.7158 | -85.177
4.7161 | -85.1765
4.7164 | -85.1757
4.7167 | -85.1752
4.717 | -85.1742
4.7174 | -85.1736
4.7177 | -85.174
4.718 | -85.1744
4.7183 | -85.1737
4.7186 | -85.1738
4.7189 | -85.1748
4.7192 | -85.1757
4.7195 | -85.1755
4.7198 | -85.1749
4.7201 | -85.1744
4.7204 | -85.1737
4.7207 | -85.1729
4.7211 | -85.1716
4.7214 | -85.1715
4.7217 | -85.1711
4.722 | -85.1708
4.7223 | -85.1705
4.7226 | -85.1701
4.7229 | -85.1702
4.7232 | -85.1704
4.7235 | -85.1703
4.7238 | -85.1704
4.7241 | -85.1699
4.7245 | -85.1693
4.7248 | -85.1699
4.7251 | -85.1709
4.7254 | -85.1714
4.7257 | -85.1711
4.726 | -85.1706
4.7263 | -85.1699
4.7266 | -85.1688
4.7269 | -85.1683
4.7272 | -85.1686
4.7275 | -85.1684
4.7279 | -85.168
4.7282 | -85.1675
4.7285 | -85.167
4.7288 | -85.1667
4.7291 | -85.1654
4.7294 | -85.1637
4.7297 | -85.1629
4.73 | -85.1622
4.7303 | -85.1611
4.7306 | -85.161
4.7309 | -85.1606
4.7312 | -85.1601
4.7316 | -85.1599
4.7319 | -85.1605
4.7322 | -85.1606
4.7325 | -85.1593
4.7328 | -85.1572
4.7331 | -85.1554
4.7334 | -85.1549
4.7337 | -85.1542
4.734 | -85.1535
4.7343 | -85.1529
4.7346 | -85.1518
4.735 | -85.1503
4.7353 | -85.149
4.7356 | -85.1486
4.7359 | -85.1482
4.7362 | -85.1481
4.7365 | -85.1482
4.7368 | -85.1475
4.7371 | -85.1468
4.7374 | -85.1462
4.7377 | -85.1457
4.738 | -85.1455
4.7383 | -85.1452
4.7387 | -85.1454
4.739 | -85.1451
4.7393 | -85.145
4.7396 | -85.1447
4.7399 | -85.1437
4.7402 | -85.1426
4.7405 | -85.1418
4.7408 | -85.1416
4.7411 | -85.1423
4.7414 | -85.1427
4.7417 | -85.1423
4.7421 | -85.1418
4.7424 | -85.1417
4.7427 | -85.1414
4.743 | -85.1408
4.7433 | -85.14
4.7436 | -85.1398
4.7439 | -85.1388
4.7442 | -85.1383
4.7445 | -85.1387
4.7448 | -85.1397
4.7451 | -85.1396
4.7455 | -85.1394
4.7458 | -85.139
4.7461 | -85.1387
4.7464 | -85.1383
4.7467 | -85.1377
4.747 | -85.1378
4.7473 | -85.1381
4.7476 | -85.1377
4.7479 | -85.1375
4.7482 | -85.1384
4.7485 | -85.1383
4.7488 | -85.1382
4.7492 | -85.138
4.7495 | -85.1381
4.7498 | -85.1382
4.7501 | -85.1386
4.7504 | -85.139
4.7507 | -85.1389
4.751 | -85.1385
4.7513 | -85.1392
4.7516 | -85.1397
4.7519 | -85.1399
4.7522 | -85.1393
4.7526 | -85.1408
4.7529 | -85.1427
4.7532 | -85.1432
4.7535 | -85.1425
4.7538 | -85.1424
4.7541 | -85.1426
4.7544 | -85.1426
4.7547 | -85.1425
4.755 | -85.1421
4.7553 | -85.1423
4.7556 | -85.1427
4.756 | -85.1428
4.7563 | -85.1427
4.7566 | -85.1426
4.7569 | -85.1426
4.7572 | -85.1437
4.7575 | -85.1439
4.7578 | -85.1437
4.7581 | -85.1435
4.7584 | -85.1434
4.7587 | -85.1434
4.759 | -85.1429
4.7593 | -85.1427
4.7597 | -85.1427
4.76 | -85.1427
4.7603 | -85.1424
4.7606 | -85.1428
4.7609 | -85.1434
4.7612 | -85.1433
4.7615 | -85.1428
4.7618 | -85.1425
4.7621 | -85.1428
4.7624 | -85.143
4.7627 | -85.1432
4.7631 | -85.1429
4.7634 | -85.1422
4.7637 | -85.142
4.764 | -85.1415
4.7643 | -85.143
4.7646 | -85.1434
4.7649 | -85.1432
4.7652 | -85.1422
4.7655 | -85.1416
4.7658 | -85.1413
4.7661 | -85.1425
4.7664 | -85.1424
4.7668 | -85.1424
4.7671 | -85.1434
4.7674 | -85.1447
4.7677 | -85.1442
4.768 | -85.1434
4.7683 | -85.1432
4.7686 | -85.1434
4.7689 | -85.1429
4.7692 | -85.1429
4.7695 | -85.1425
4.7698 | -85.1442
4.7702 | -85.1473
4.7705 | -85.1485
4.7708 | -85.1484
4.7711 | -85.1497
4.7714 | -85.1505
4.7717 | -85.1506
4.772 | -85.1509
4.7723 | -85.1518
4.7726 | -85.1535
4.7729 | -85.1552
4.7732 | -85.1558
4.7736 | -85.1559
4.7739 | -85.156
4.7742 | -85.1565
4.7745 | -85.1567
4.7748 | -85.1551
4.7751 | -85.1554
4.7754 | -85.1557
4.7757 | -85.1556
4.776 | -85.1557
4.7763 | -85.156
4.7766 | -85.1565
4.7769 | -85.1567
4.7773 | -85.1556
4.7776 | -85.1547
4.7779 | -85.1544
4.7782 | -85.1546
4.7785 | -85.1554
4.7788 | -85.1554
4.7791 | -85.1559
4.7794 | -85.1556
4.7797 | -85.1564
4.78 | -85.1572
4.7803 | -85.1575
4.7807 | -85.1582
4.781 | -85.1581
4.7813 | -85.1579
4.7816 | -85.1573
4.7819 | -85.1569
4.7822 | -85.1574
4.7825 | -85.1572
4.7828 | -85.1578
4.7831 | -85.1586
4.7834 | -85.1598
4.7837 | -85.1612
4.784 | -85.1623
4.7844 | -85.1628
4.7847 | -85.1635
4.785 | -85.1638
4.7853 | -85.1641
4.7856 | -85.1644
4.7859 | -85.1648
4.7862 | -85.1658
4.7865 | -85.1669
4.7868 | -85.1676
4.7871 | -85.1676
4.7874 | -85.1667
4.7878 | -85.1663
4.7881 | -85.1664
4.7884 | -85.1664
4.7887 | -85.1665
4.789 | -85.167
4.7893 | -85.1678
4.7896 | -85.1689
4.7899 | -85.1696
4.7902 | -85.17
4.7905 | -85.1697
4.7908 | -85.17
4.7911 | -85.1705
4.7915 | -85.1711
4.7918 | -85.1719
4.7921 | -85.1727
4.7924 | -85.1734
4.7927 | -85.1751
4.793 | -85.1749
4.7933 | -85.1755
4.7936 | -85.1763
4.7939 | -85.1774
4.7942 | -85.1779
4.7945 | -85.1778
4.7949 | -85.1784
4.7952 | -85.179
4.7955 | -85.1792
4.7958 | -85.1798
4.7961 | -85.1806
4.7964 | -85.1813
4.7967 | -85.1823
4.797 | -85.1833
4.7973 | -85.184
4.7976 | -85.185
4.7979 | -85.1859
4.7983 | -85.1869
4.7986 | -85.1883
4.7989 | -85.1897
4.7992 | -85.1906
4.7995 | -85.1914
4.7998 | -85.1925
4.8001 | -85.1944
4.8004 | -85.1955
4.8007 | -85.1965
4.801 | -85.1972
4.8013 | -85.1975
4.8016 | -85.1994
4.802 | -85.2007
4.8023 | -85.202
4.8026 | -85.2026
4.8029 | -85.2034
4.8032 | -85.2042
4.8035 | -85.2056
4.8038 | -85.207
4.8041 | -85.2074
4.8044 | -85.2079
4.8047 | -85.209
4.805 | -85.2108
4.8054 | -85.2115
4.8057 | -85.2124
4.806 | -85.2132
4.8063 | -85.2142
4.8066 | -85.2148
4.8069 | -85.2148
4.8072 | -85.2156
4.8075 | -85.2176
4.8078 | -85.2189
4.8081 | -85.22
4.8084 | -85.2216
4.8087 | -85.2237
4.8091 | -85.2256
4.8094 | -85.2268
4.8097 | -85.2277
4.81 | -85.2294
4.8103 | -85.2311
4.8106 | -85.2331
4.8109 | -85.2347
4.8112 | -85.2353
4.8115 | -85.2355
4.8118 | -85.2373
4.8121 | -85.2391
4.8125 | -85.2402
4.8128 | -85.2413
4.8131 | -85.2428
4.8134 | -85.2442
4.8137 | -85.2451
4.814 | -85.2472
4.8143 | -85.2492
4.8146 | -85.2508
4.8149 | -85.2518
4.8152 | -85.2526
4.8155 | -85.2527
4.8159 | -85.253
4.8162 | -85.254
4.8165 | -85.2552
4.8168 | -85.2561
4.8171 | -85.257
4.8174 | -85.2578
4.8177 | -85.2595
4.818 | -85.2613
4.8183 | -85.2637
4.8186 | -85.2662
4.8189 | -85.2684
4.8192 | -85.2696
4.8196 | -85.2709
4.8199 | -85.2719
4.8202 | -85.2727
4.8205 | -85.2735
4.8208 | -85.274
4.8211 | -85.2749
4.8214 | -85.2767
4.8217 | -85.2782
4.822 | -85.2793
4.8223 | -85.2812
4.8226 | -85.2831
4.823 | -85.2846
4.8233 | -85.2861
4.8236 | -85.2876
4.8239 | -85.2895
4.8242 | -85.2909
4.8245 | -85.2922
4.8248 | -85.2941
4.8251 | -85.2963
4.8254 | -85.2981
4.8257 | -85.2991
4.826 | -85.3009
4.8263 | -85.3017
4.8267 | -85.3032
4.827 | -85.3041
4.8273 | -85.3056
4.8276 | -85.3069
4.8279 | -85.3081
4.8282 | -85.3096
4.8285 | -85.3103
4.8288 | -85.3123
4.8291 | -85.3138
4.8294 | -85.3157
4.8297 | -85.3163
4.8301 | -85.3161
4.8304 | -85.3185
4.8307 | -85.3203
4.831 | -85.3206
4.8313 | -85.3219
4.8316 | -85.3224
4.8319 | -85.3236
4.8322 | -85.3248
4.8325 | -85.3262
4.8328 | -85.3276
4.8331 | -85.3291
4.8335 | -85.3311
4.8338 | -85.3345
4.8341 | -85.3367
4.8344 | -85.338
4.8347 | -85.3389
4.835 | -85.3404
4.8353 | -85.3416
4.8356 | -85.3431
4.8359 | -85.3443
4.8362 | -85.3456
4.8365 | -85.347
4.8368 | -85.3481
4.8372 | -85.3504
4.8375 | -85.353
4.8378 | -85.355
4.8381 | -85.3574
4.8384 | -85.3589
4.8387 | -85.3601
4.839 | -85.3621
4.8393 | -85.3655
4.8396 | -85.3672
4.8399 | -85.3679
4.8402 | -85.3698
4.8406 | -85.372
4.8409 | -85.3737
4.8412 | -85.3754
4.8415 | -85.3761
4.8418 | -85.3776
4.8421 | -85.3795
4.8424 | -85.3807
4.8427 | -85.3808
4.843 | -85.381
4.8433 | -85.382
4.8436 | -85.3836
4.8439 | -85.3846
4.8443 | -85.3852
4.8446 | -85.3866
4.8449 | -85.3881
4.8452 | -85.3902
4.8455 | -85.3916
4.8458 | -85.3924
4.8461 | -85.393
4.8464 | -85.3934
4.8467 | -85.3941
4.847 | -85.3956
4.8473 | -85.399
4.8477 | -85.4012
4.848 | -85.4033
4.8483 | -85.4054
4.8486 | -85.4066
4.8489 | -85.4078
4.8492 | -85.409
4.8495 | -85.4093
4.8498 | -85.4109
4.8501 | -85.4122
4.8504 | -85.4132
4.8507 | -85.4139
4.851 | -85.4149
4.8514 | -85.4181
4.8517 | -85.4211
4.852 | -85.4229
4.8523 | -85.425
4.8526 | -85.4263
4.8529 | -85.4278
4.8532 | -85.4295
4.8535 | -85.432
4.8538 | -85.4332
4.8541 | -85.4349
4.8544 | -85.4363
4.8548 | -85.4377
4.8551 | -85.4397
4.8554 | -85.4413
4.8557 | -85.4436
4.856 | -85.4458
4.8563 | -85.4473
4.8566 | -85.4494
4.8569 | -85.4522
4.8572 | -85.4549
4.8575 | -85.4567
4.8578 | -85.4586
4.8582 | -85.4611
4.8585 | -85.4643
4.8588 | -85.4664
4.8591 | -85.468
4.8594 | -85.4696
4.8597 | -85.4708
4.86 | -85.4728
4.8603 | -85.4753
4.8606 | -85.4767
4.8609 | -85.4794
4.8612 | -85.4834
4.8615 | -85.4846
4.8619 | -85.4859
4.8622 | -85.4874
4.8625 | -85.4889
4.8628 | -85.4905
4.8631 | -85.4923
4.8634 | -85.4943
4.8637 | -85.4963
4.864 | -85.4986
4.8643 | -85.5002
4.8646 | -85.5009
4.8649 | -85.5027
4.8653 | -85.5046
4.8656 | -85.5056
4.8659 | -85.5062
4.8662 | -85.5076
4.8665 | -85.5085
4.8668 | -85.5108
4.8671 | -85.5119
4.8674 | -85.5131
4.8677 | -85.515
4.868 | -85.5166
4.8683 | -85.5183
4.8686 | -85.52
4.869 | -85.5213
4.8693 | -85.5222
4.8696 | -85.5239
4.8699 | -85.5256
4.8702 | -85.5273
4.8705 | -85.5287
4.8708 | -85.53
4.8711 | -85.5315
4.8714 | -85.5329
4.8717 | -85.5352
4.872 | -85.5374
4.8724 | -85.5389
4.8727 | -85.5404
4.873 | -85.5419
4.8733 | -85.5428
4.8736 | -85.5444
4.8739 | -85.5461
4.8742 | -85.5481
4.8745 | -85.5492
4.8748 | -85.5509
4.8751 | -85.5523
4.8754 | -85.5543
4.8758 | -85.5563
4.8761 | -85.5581
4.8764 | -85.5593
4.8767 | -85.5606
4.877 | -85.5626
4.8773 | -85.5646
4.8776 | -85.5663
4.8779 | -85.5676
4.8782 | -85.57
4.8785 | -85.5717
4.8788 | -85.5736
4.8791 | -85.5736
4.8795 | -85.5717
4.8798 | -85.5715
4.8801 | -85.5729
4.8804 | -85.5742
4.8807 | -85.5765
4.881 | -85.579
4.8813 | -85.5802
4.8816 | -85.5822
4.8819 | -85.5837
4.8822 | -85.5855
4.8825 | -85.5869
4.8829 | -85.5881
4.8832 | -85.5896
4.8835 | -85.5908
4.8838 | -85.5924
4.8841 | -85.5936
4.8844 | -85.5946
4.8847 | -85.5964
4.885 | -85.5989
4.8853 | -85.6021
4.8856 | -85.6056
4.8859 | -85.6069
4.8863 | -85.6081
4.8866 | -85.6099
4.8869 | -85.611
4.8872 | -85.6118
4.8875 | -85.6118
4.8878 | -85.6132
4.8881 | -85.6149
4.8884 | -85.6168
4.8887 | -85.6189
4.889 | -85.6193
4.8893 | -85.6188
4.8896 | -85.6201
4.89 | -85.6217
4.8903 | -85.6229
4.8906 | -85.6244
4.8909 | -85.6256
4.8912 | -85.6275
4.8915 | -85.6288
4.8918 | -85.6301
4.8921 | -85.6319
4.8924 | -85.6335
4.8927 | -85.6348
4.893 | -85.6367
4.8934 | -85.6378
4.8937 | -85.6388
4.894 | -85.6401
4.8943 | -85.6417
4.8946 | -85.6431
4.8949 | -85.644
4.8952 | -85.6454
4.8955 | -85.6464
4.8958 | -85.6477
4.8961 | -85.6496
4.8964 | -85.651
4.8967 | -85.652
4.8971 | -85.6533
4.8974 | -85.6547
4.8977 | -85.6561
4.898 | -85.6568
4.8983 | -85.6573
4.8986 | -85.6583
4.8989 | -85.6596
4.8992 | -85.6608
4.8995 | -85.6625
4.8998 | -85.6639
4.9001 | -85.6648
4.9005 | -85.6657
4.9008 | -85.6666
4.9011 | -85.6679
4.9014 | -85.6691
4.9017 | -85.67
4.902 | -85.6711
4.9023 | -85.6724
4.9026 | -85.6738
4.9029 | -85.6755
4.9032 | -85.6768
4.9035 | -85.6779
4.9039 | -85.679
4.9042 | -85.6802
4.9045 | -85.6828
4.9048 | -85.685
4.9051 | -85.6864
4.9054 | -85.6883
4.9057 | -85.6895
4.906 | -85.6916
4.9063 | -85.6927
4.9066 | -85.6942
4.9069 | -85.6955
4.9072 | -85.6975
4.9076 | -85.6998
4.9079 | -85.7009
4.9082 | -85.7014
4.9085 | -85.7018
4.9088 | -85.703
4.9091 | -85.7043
4.9094 | -85.7054
4.9097 | -85.7065
4.91 | -85.7078
4.9103 | -85.7085
4.9106 | -85.7096
4.911 | -85.7104
4.9113 | -85.7103
4.9116 | -85.712
4.9119 | -85.7132
4.9122 | -85.7144
4.9125 | -85.7154
4.9128 | -85.7162
4.9131 | -85.7183
4.9134 | -85.7197
4.9137 | -85.721
4.914 | -85.7221
4.9143 | -85.7226
4.9147 | -85.7236
4.915 | -85.7246
4.9153 | -85.7254
4.9156 | -85.7265
4.9159 | -85.7281
4.9162 | -85.7302
4.9165 | -85.732
4.9168 | -85.7331
4.9171 | -85.7342
4.9174 | -85.7345
4.9177 | -85.7356
4.9181 | -85.7384
4.9184 | -85.7403
4.9187 | -85.7415
4.919 | -85.7431
4.9193 | -85.744
4.9196 | -85.7445
4.9199 | -85.7457
4.9202 | -85.7469
4.9205 | -85.7486
4.9208 | -85.7495
4.9211 | -85.7506
4.9214 | -85.7529
4.9218 | -85.7562
4.9221 | -85.7579
4.9224 | -85.7588
4.9227 | -85.7598
4.923 | -85.7611
4.9233 | -85.7623
4.9236 | -85.764
4.9239 | -85.7654
4.9242 | -85.77
4.9245 | -85.7749
4.9248 | -85.7758
4.9252 | -85.7764
4.9255 | -85.7774
4.9258 | -85.7789
4.9261 | -85.7811
4.9264 | -85.7803
4.9267 | -85.7823
4.927 | -85.7849
4.9273 | -85.7886
4.9276 | -85.7908
4.9279 | -85.792
4.9282 | -85.793
4.9286 | -85.794
4.9289 | -85.7943
4.9292 | -85.7943
4.9295 | -85.7947
4.9298 | -85.7955
4.9301 | -85.7968
4.9304 | -85.7979
4.9307 | -85.7973
4.931 | -85.7966
4.9313 | -85.7964
4.9316 | -85.797
4.9319 | -85.7979
4.9323 | -85.7989
4.9326 | -85.7994
4.9329 | -85.7994
4.9332 | -85.8007
4.9335 | -85.8018
4.9338 | -85.8027
4.9341 | -85.8042
4.9344 | -85.8054
4.9347 | -85.8057
4.935 | -85.807
4.9353 | -85.8082
4.9357 | -85.8097
4.936 | -85.8104
4.9363 | -85.8113
4.9366 | -85.8128
4.9369 | -85.8143
4.9372 | -85.814
4.9375 | -85.813
4.9378 | -85.8138
4.9381 | -85.8146
4.9384 | -85.8152
4.9387 | -85.8162
4.939 | -85.8169
4.9394 | -85.8181
4.9397 | -85.8185
4.94 | -85.819
4.9403 | -85.8193
4.9406 | -85.8192
4.9409 | -85.8189
4.9412 | -85.8194
4.9415 | -85.82
4.9418 | -85.8202
4.9421 | -85.8206
4.9424 | -85.8205
4.9428 | -85.82
4.9431 | -85.8199
4.9434 | -85.8199
4.9437 | -85.8206
4.944 | -85.8208
4.9443 | -85.8207
4.9446 | -85.8213
4.9449 | -85.822
4.9452 | -85.8225
4.9455 | -85.8228
4.9458 | -85.8237
4.9462 | -85.8246
4.9465 | -85.8257
4.9468 | -85.8267
4.9471 | -85.8271
4.9474 | -85.8283
4.9477 | -85.8291
4.948 | -85.8295
4.9483 | -85.8307
4.9486 | -85.8327
4.9489 | -85.8333
4.9492 | -85.8338
4.9495 | -85.8344
4.9499 | -85.8347
4.9502 | -85.835
4.9505 | -85.8349
4.9508 | -85.835
4.9511 | -85.8355
4.9514 | -85.8361
4.9517 | -85.8365
4.952 | -85.837
4.9523 | -85.8384
4.9526 | -85.8399
4.9529 | -85.8409
4.9533 | -85.8427
4.9536 | -85.8427
4.9539 | -85.8422
4.9542 | -85.8429
4.9545 | -85.844
4.9548 | -85.8443
4.9551 | -85.8446
4.9554 | -85.846
4.9557 | -85.847
4.956 | -85.8477
4.9563 | -85.8484
4.9566 | -85.8493
4.957 | -85.8505
4.9573 | -85.8518
4.9576 | -85.8525
4.9579 | -85.8529
4.9582 | -85.8541
4.9585 | -85.8548
4.9588 | -85.8555
4.9591 | -85.8563
4.9594 | -85.8569
4.9597 | -85.8578
4.96 | -85.8583
4.9604 | -85.8588
4.9607 | -85.8594
4.961 | -85.8606
4.9613 | -85.8625
4.9616 | -85.8634
4.9619 | -85.8644
4.9622 | -85.8647
4.9625 | -85.8653
4.9628 | -85.8662
4.9631 | -85.8665
4.9634 | -85.8666
4.9638 | -85.8656
4.9641 | -85.8647
4.9644 | -85.8647
4.9647 | -85.8651
4.965 | -85.8668
4.9653 | -85.8684
4.9656 | -85.8694
4.9659 | -85.8694
4.9662 | -85.8695
4.9665 | -85.8699
4.9668 | -85.8708
4.9671 | -85.8716
4.9675 | -85.8723
4.9678 | -85.8741
4.9681 | -85.876
4.9684 | -85.8773
4.9687 | -85.8781
4.969 | -85.8785
4.9693 | -85.8793
4.9696 | -85.8806
4.9699 | -85.882
4.9702 | -85.8822
4.9705 | -85.8821
4.9709 | -85.8833
4.9712 | -85.8846
4.9715 | -85.8854
4.9718 | -85.8865
4.9721 | -85.8874
4.9724 | -85.8884
4.9727 | -85.8897
4.973 | -85.8899
4.9733 | -85.891
4.9736 | -85.8917
4.9739 | -85.8922
4.9742 | -85.8929
4.9746 | -85.8947
4.9749 | -85.8959
4.9752 | -85.8958
4.9755 | -85.8967
4.9758 | -85.8968
4.9761 | -85.898
4.9764 | -85.8993
4.9767 | -85.9002
4.977 | -85.9016
4.9773 | -85.9019
4.9776 | -85.9027
4.978 | -85.9036
4.9783 | -85.9041
4.9786 | -85.9044
4.9789 | -85.9046
4.9792 | -85.9051
4.9795 | -85.9055
4.9798 | -85.9069
4.9801 | -85.9076
4.9804 | -85.9085
4.9807 | -85.909
4.981 | -85.91
4.9813 | -85.9103
4.9817 | -85.9098
4.982 | -85.9104
4.9823 | -85.9117
4.9826 | -85.9139
4.9829 | -85.9144
4.9832 | -85.9151
4.9835 | -85.9165
4.9838 | -85.9171
4.9841 | -85.9191
4.9844 | -85.9203
4.9847 | -85.9214
4.9851 | -85.9229
4.9854 | -85.9221
4.9857 | -85.9223
4.986 | -85.9231
4.9863 | -85.9234
4.9866 | -85.9237
4.9869 | -85.9242
4.9872 | -85.925
4.9875 | -85.926
4.9878 | -85.9261
4.9881 | -85.9263
4.9885 | -85.927
4.9888 | -85.9284
4.9891 | -85.929
4.9894 | -85.9302
4.9897 | -85.9313
4.99 | -85.9327
4.9903 | -85.9345
4.9906 | -85.9367
4.9909 | -85.9373
4.9912 | -85.9388
4.9915 | -85.9394
4.9918 | -85.9401
4.9922 | -85.9407
4.9925 | -85.9414
4.9928 | -85.9416
4.9931 | -85.9421
4.9934 | -85.9426
4.9937 | -85.9429
4.994 | -85.944
4.9943 | -85.9463
4.9946 | -85.9473
4.9949 | -85.9476
4.9952 | -85.9479
4.9956 | -85.9495
4.9959 | -85.9507
4.9962 | -85.9511
4.9965 | -85.9517
4.9968 | -85.9519
4.9971 | -85.9513
4.9974 | -85.951
4.9977 | -85.9524
4.998 | -85.9538
4.9983 | -85.9548
4.9986 | -85.9554
4.9989 | -85.9566
4.9993 | -85.9575
4.9996 | -85.9597
4.9999 | -85.962
5.0002 | -85.9638
5.0005 | -85.9645
5.0008 | -85.9648
5.0011 | -85.965
5.0014 | -85.9652
5.0017 | -85.967
5.002 | -85.9718
5.0023 | -85.9732
5.0027 | -85.9727
5.003 | -85.9729
5.0033 | -85.9742
5.0036 | -85.975
5.0039 | -85.9749
5.0042 | -85.9754
5.0045 | -85.9762
5.0048 | -85.9768
5.0051 | -85.9777
5.0054 | -85.9785
5.0057 | -85.9777
5.0061 | -85.978
5.0064 | -85.979
5.0067 | -85.9803
5.007 | -85.9808
5.0073 | -85.9814
5.0076 | -85.9826
5.0079 | -85.9843
5.0082 | -85.9853
5.0085 | -85.9861
5.0088 | -85.9867
5.0091 | -85.9873
5.0094 | -85.9891
5.0098 | -85.9938
5.0101 | -85.997
5.0104 | -85.998
5.0107 | -85.9986
5.011 | -85.999
5.0113 | -86.0002
5.0116 | -86.0012
5.0119 | -86.0017
5.0122 | -86.0021
5.0125 | -86.0033
5.0128 | -86.0046
5.0132 | -86.005
5.0135 | -86.0052
5.0138 | -86.0065
5.0141 | -86.0075
5.0144 | -86.0081
5.0147 | -86.0092
5.015 | -86.0096
5.0153 | -86.0106
5.0156 | -86.011
5.0159 | -86.0109
5.0162 | -86.0117
5.0165 | -86.0128
5.0169 | -86.0138
5.0172 | -86.0148
5.0175 | -86.0154
5.0178 | -86.0162
5.0181 | -86.0176
5.0184 | -86.0194
5.0187 | -86.0217
5.019 | -86.0243
5.0193 | -86.0254
5.0196 | -86.0257
5.0199 | -86.0261
5.0203 | -86.0265
5.0206 | -86.027
5.0209 | -86.0283
5.0212 | -86.0285
5.0215 | -86.029
5.0218 | -86.0302
5.0221 | -86.0322
5.0224 | -86.0343
5.0227 | -86.0355
5.023 | -86.0363
5.0233 | -86.0364
5.0237 | -86.0356
5.024 | -86.0361
5.0243 | -86.0376
5.0246 | -86.039
5.0249 | -86.0404
5.0252 | -86.0415
5.0255 | -86.0424
5.0258 | -86.0432
5.0261 | -86.0443
5.0264 | -86.0454
5.0267 | -86.0465
5.027 | -86.0476
5.0274 | -86.0476
5.0277 | -86.0481
5.028 | -86.0487
5.0283 | -86.0494
5.0286 | -86.0503
5.0289 | -86.0513
5.0292 | -86.0519
5.0295 | -86.0528
5.0298 | -86.0535
5.0301 | -86.0546
5.0304 | -86.0556
5.0308 | -86.0561
5.0311 | -86.0561
5.0314 | -86.0561
5.0317 | -86.0566
5.032 | -86.0576
5.0323 | -86.0577
5.0326 | -86.0591
5.0329 | -86.0598
5.0332 | -86.0609
5.0335 | -86.0622
5.0338 | -86.0641
5.0342 | -86.0645
5.0345 | -86.0653
5.0348 | -86.0675
5.0351 | -86.0683
5.0354 | -86.0692
5.0357 | -86.0702
5.036 | -86.0718
5.0363 | -86.0727
5.0366 | -86.0735
5.0369 | -86.0747
5.0372 | -86.0762
5.0375 | -86.0774
5.0379 | -86.0786
5.0382 | -86.0792
5.0385 | -86.0819
5.0388 | -86.083
5.0391 | -86.0837
5.0394 | -86.0846
5.0397 | -86.0857
5.04 | -86.0875
5.0403 | -86.0891
5.0406 | -86.0907
5.0409 | -86.0926
5.0413 | -86.0951
5.0416 | -86.0974
5.0419 | -86.0984
5.0422 | -86.1002
5.0425 | -86.101
5.0428 | -86.1024
5.0431 | -86.1033
5.0434 | -86.1038
5.0437 | -86.1046
5.044 | -86.106
5.0443 | -86.1082
5.0446 | -86.1093
5.045 | -86.1105
5.0453 | -86.1114
5.0456 | -86.1133
5.0459 | -86.1152
5.0462 | -86.1161
5.0465 | -86.1176
5.0468 | -86.1181
5.0471 | -86.1191
5.0474 | -86.1205
5.0477 | -86.1216
5.048 | -86.1223
5.0484 | -86.1232
5.0487 | -86.1243
5.049 | -86.1258
5.0493 | -86.1277
5.0496 | -86.1292
5.0499 | -86.1305
5.0502 | -86.1317
5.0505 | -86.1328
5.0508 | -86.1336
5.0511 | -86.1354
5.0514 | -86.1364
5.0517 | -86.137
5.0521 | -86.1377
5.0524 | -86.1387
5.0527 | -86.1399
5.053 | -86.1413
5.0533 | -86.1427
5.0536 | -86.1441
5.0539 | -86.1454
5.0542 | -86.1471
5.0545 | -86.1482
5.0548 | -86.1496
5.0551 | -86.1507
5.0555 | -86.152
5.0558 | -86.153
5.0561 | -86.1533
5.0564 | -86.1541
5.0567 | -86.155
5.057 | -86.1558
5.0573 | -86.157
5.0576 | -86.1583
5.0579 | -86.1594
5.0582 | -86.1607
5.0585 | -86.1617
5.0589 | -86.1624
5.0592 | -86.1633
5.0595 | -86.1644
5.0598 | -86.1653
5.0601 | -86.1674
5.0604 | -86.1684
5.0607 | -86.1698
5.061 | -86.1713
5.0613 | -86.1726
5.0616 | -86.1738
5.0619 | -86.1739
5.0622 | -86.1743
5.0626 | -86.1757
5.0629 | -86.1766
5.0632 | -86.1776
5.0635 | -86.1784
5.0638 | -86.179
5.0641 | -86.1803
5.0644 | -86.1811
5.0647 | -86.1823
5.065 | -86.1833
5.0653 | -86.1844
5.0656 | -86.1853
5.066 | -86.1865
5.0663 | -86.188
5.0666 | -86.1902
5.0669 | -86.1919
5.0672 | -86.1929
5.0675 | -86.194
5.0678 | -86.195
5.0681 | -86.196
5.0684 | -86.1974
5.0687 | -86.1985
5.069 | -86.1994
5.0693 | -86.2017
5.0697 | -86.2044
5.07 | -86.2063
5.0703 | -86.2084
5.0706 | -86.2106
5.0709 | -86.212
5.0712 | -86.2132
5.0715 | -86.2146
5.0718 | -86.2162
5.0721 | -86.2176
5.0724 | -86.2187
5.0727 | -86.2204
5.0731 | -86.2227
5.0734 | -86.2263
5.0737 | -86.2293
5.074 | -86.2301
5.0743 | -86.2317
5.0746 | -86.2334
5.0749 | -86.2344
5.0752 | -86.2364
5.0755 | -86.2382
5.0758 | -86.24
5.0761 | -86.2413
5.0765 | -86.2425
5.0768 | -86.2438
5.0771 | -86.2452
5.0774 | -86.2469
5.0777 | -86.2458
5.078 | -86.2458
5.0783 | -86.247
5.0786 | -86.2487
5.0789 | -86.2501
5.0792 | -86.2515
5.0795 | -86.2529
5.0798 | -86.2541
5.0802 | -86.2554
5.0805 | -86.2562
5.0808 | -86.2576
5.0811 | -86.2598
5.0814 | -86.2618
5.0817 | -86.2631
5.082 | -86.2638
5.0823 | -86.2653
5.0826 | -86.2668
5.0829 | -86.2678
5.0832 | -86.2681
5.0836 | -86.2689
5.0839 | -86.2712
5.0842 | -86.2728
5.0845 | -86.2744
5.0848 | -86.2737
5.0851 | -86.2727
5.0854 | -86.2724
5.0857 | -86.2736
5.086 | -86.2748
5.0863 | -86.2759
5.0866 | -86.2774
5.0869 | -86.2792
5.0873 | -86.2808
5.0876 | -86.2821
5.0879 | -86.2834
5.0882 | -86.2854
5.0885 | -86.2869
5.0888 | -86.2882
5.0891 | -86.2904
5.0894 | -86.2917
5.0897 | -86.293
5.09 | -86.2936
5.0903 | -86.2949
5.0907 | -86.2964
5.091 | -86.2977
5.0913 | -86.2989
5.0916 | -86.2996
5.0919 | -86.3001
5.0922 | -86.301
5.0925 | -86.3024
5.0928 | -86.3042
5.0931 | -86.3049
5.0934 | -86.3054
5.0937 | -86.3069
5.0941 | -86.3089
5.0944 | -86.3107
5.0947 | -86.3117
5.095 | -86.3134
5.0953 | -86.3163
5.0956 | -86.3179
5.0959 | -86.3192
5.0962 | -86.3203
5.0965 | -86.3218
5.0968 | -86.3228
5.0971 | -86.3238
5.0974 | -86.3254
5.0978 | -86.3272
5.0981 | -86.3288
5.0984 | -86.3302
5.0987 | -86.3317
5.099 | -86.3335
5.0993 | -86.3351
5.0996 | -86.3366
5.0999 | -86.338
5.1002 | -86.3393
5.1005 | -86.3413
5.1008 | -86.3431
5.1012 | -86.3447
5.1015 | -86.3467
5.1018 | -86.3482
5.1021 | -86.3494
5.1024 | -86.3507
5.1027 | -86.3525
5.103 | -86.3543
5.1033 | -86.3558
5.1036 | -86.3569
5.1039 | -86.3582
5.1042 | -86.3588
5.1045 | -86.3603
5.1049 | -86.3611
5.1052 | -86.3622
5.1055 | -86.3637
5.1058 | -86.3657
5.1061 | -86.3685
5.1064 | -86.3705
5.1067 | -86.3722
5.107 | -86.3738
5.1073 | -86.3749
5.1076 | -86.3753
5.1079 | -86.3777
5.1083 | -86.38
5.1086 | -86.3817
5.1089 | -86.3832
5.1092 | -86.3846
5.1095 | -86.3865
5.1098 | -86.3878
5.1101 | -86.3882
5.1104 | -86.3896
5.1107 | -86.392
5.111 | -86.3949
5.1113 | -86.3968
5.1117 | -86.3983
5.112 | -86.4002
5.1123 | -86.4009
5.1126 | -86.4028
5.1129 | -86.4051
5.1132 | -86.4064
5.1135 | -86.4078
5.1138 | -86.4093
5.1141 | -86.4114
5.1144 | -86.4128
5.1147 | -86.4139
5.115 | -86.4139
5.1154 | -86.4146
5.1157 | -86.416
5.116 | -86.4176
5.1163 | -86.4185
5.1166 | -86.4198
5.1169 | -86.4216
5.1172 | -86.4239
5.1175 | -86.425
5.1178 | -86.4268
5.1181 | -86.4301
5.1184 | -86.4327
5.1188 | -86.4338
5.1191 | -86.4351
5.1194 | -86.4365
5.1197 | -86.438
5.12 | -86.4395
5.1203 | -86.4409
5.1206 | -86.4425
5.1209 | -86.4458
5.1212 | -86.4489
5.1215 | -86.451
5.1218 | -86.4527
5.1221 | -86.4546
5.1225 | -86.4562
5.1228 | -86.4582
5.1231 | -86.4591
5.1234 | -86.461
5.1237 | -86.4626
5.124 | -86.4638
5.1243 | -86.4655
5.1246 | -86.4664
5.1249 | -86.4675
5.1252 | -86.4703
5.1255 | -86.4732
5.1259 | -86.4753
5.1262 | -86.4765
5.1265 | -86.4786
5.1268 | -86.4809
5.1271 | -86.4814
5.1274 | -86.4806
5.1277 | -86.4804
5.128 | -86.4812
5.1283 | -86.4824
5.1286 | -86.484
5.1289 | -86.4846
5.1292 | -86.4858
5.1296 | -86.4872
5.1299 | -86.4889
5.1302 | -86.4907
5.1305 | -86.4921
5.1308 | -86.4931
5.1311 | -86.4938
5.1314 | -86.4948
5.1317 | -86.4967
5.132 | -86.4995
5.1323 | -86.5019
5.1326 | -86.5038
5.133 | -86.5055
5.1333 | -86.5073
5.1336 | -86.5083
5.1339 | -86.5091
5.1342 | -86.5106
5.1345 | -86.5117
5.1348 | -86.5135
5.1351 | -86.5159
5.1354 | -86.5178
5.1357 | -86.5192
5.136 | -86.5205
5.1364 | -86.5218
5.1367 | -86.5232
5.137 | -86.5242
5.1373 | -86.5245
5.1376 | -86.5256
5.1379 | -86.5278
5.1382 | -86.5289
5.1385 | -86.5301
5.1388 | -86.5316
5.1391 | -86.5342
5.1394 | -86.5362
5.1397 | -86.5357
5.1401 | -86.5356
5.1404 | -86.5369
5.1407 | -86.5383
5.141 | -86.54
5.1413 | -86.5412
5.1416 | -86.5424
5.1419 | -86.543
5.1422 | -86.5434
5.1425 | -86.5448
5.1428 | -86.546
5.1431 | -86.5469
5.1435 | -86.548
5.1438 | -86.5495
5.1441 | -86.5508
5.1444 | -86.5522
5.1447 | -86.5531
5.145 | -86.5542
5.1453 | -86.5556
5.1456 | -86.5567
5.1459 | -86.5577
5.1462 | -86.5593
5.1465 | -86.561
5.1468 | -86.5623
5.1472 | -86.5636
5.1475 | -86.5649
5.1478 | -86.5661
5.1481 | -86.5674
5.1484 | -86.5687
5.1487 | -86.5692
5.149 | -86.5703
5.1493 | -86.5711
5.1496 | -86.5712
5.1499 | -86.5732
5.1502 | -86.5752
5.1506 | -86.5754
5.1509 | -86.5732
5.1512 | -86.5735
5.1515 | -86.5749
5.1518 | -86.5763
5.1521 | -86.5772
5.1524 | -86.5788
5.1527 | -86.5812
5.153 | -86.5832
5.1533 | -86.5847
5.1536 | -86.585
5.154 | -86.5852
5.1543 | -86.5862
5.1546 | -86.5872
5.1549 | -86.5887
5.1552 | -86.5907
5.1555 | -86.592
5.1558 | -86.5935
5.1561 | -86.5954
5.1564 | -86.5967
5.1567 | -86.5974
5.157 | -86.5985
5.1573 | -86.6004
5.1577 | -86.602
5.158 | -86.6016
5.1583 | -86.6027
5.1586 | -86.6046
5.1589 | -86.6063
5.1592 | -86.6069
5.1595 | -86.606
5.1598 | -86.6056
5.1601 | -86.6068
5.1604 | -86.6078
5.1607 | -86.6087
5.1611 | -86.6096
5.1614 | -86.6115
5.1617 | -86.6133
5.162 | -86.6156
5.1623 | -86.6179
5.1626 | -86.6181
5.1629 | -86.6171
5.1632 | -86.6182
5.1635 | -86.62
5.1638 | -86.6218
5.1641 | -86.6226
5.1645 | -86.6231
5.1648 | -86.6223
5.1651 | -86.6227
5.1654 | -86.6227
5.1657 | -86.6227
5.166 | -86.6243
5.1663 | -86.6245
5.1666 | -86.6252
5.1669 | -86.6255
5.1672 | -86.627
5.1675 | -86.6286
5.1678 | -86.6302
5.1682 | -86.6309
5.1685 | -86.6316
5.1688 | -86.6324
5.1691 | -86.6333
5.1694 | -86.634
5.1697 | -86.6343
5.17 | -86.6349
5.1703 | -86.6356
5.1706 | -86.6365
5.1709 | -86.6373
5.1712 | -86.6383
5.1716 | -86.6396
5.1719 | -86.6401
5.1722 | -86.64
5.1725 | -86.64
5.1728 | -86.6412
5.1731 | -86.6431
5.1734 | -86.6447
5.1737 | -86.6449
5.174 | -86.6463
5.1743 | -86.647
5.1746 | -86.6475
5.1749 | -86.6487
5.1753 | -86.6497
5.1756 | -86.6518
5.1759 | -86.6517
5.1762 | -86.6519
5.1765 | -86.6518
5.1768 | -86.6529
5.1771 | -86.6536
5.1774 | -86.6551
5.1777 | -86.658
5.178 | -86.6592
5.1783 | -86.6602
5.1787 | -86.6624
5.179 | -86.6645
5.1793 | -86.6669
5.1796 | -86.6678
5.1799 | -86.6685
5.1802 | -86.6692
5.1805 | -86.669
5.1808 | -86.6694
5.1811 | -86.67
5.1814 | -86.6708
5.1817 | -86.6707
5.182 | -86.6719
5.1824 | -86.6729
5.1827 | -86.6725
5.183 | -86.6733
5.1833 | -86.6735
5.1836 | -86.6742
5.1839 | -86.6745
5.1842 | -86.6753
5.1845 | -86.6768
5.1848 | -86.6777
5.1851 | -86.6783
5.1854 | -86.6789
5.1858 | -86.6797
5.1861 | -86.6803
5.1864 | -86.6804
5.1867 | -86.6814
5.187 | -86.6821
5.1873 | -86.6819
5.1876 | -86.6814
5.1879 | -86.6816
5.1882 | -86.6831
5.1885 | -86.6837
5.1888 | -86.6846
5.1892 | -86.6852
5.1895 | -86.6852
5.1898 | -86.6867
5.1901 | -86.6874
5.1904 | -86.6875
5.1907 | -86.6881
5.191 | -86.6905
5.1913 | -86.6925
5.1916 | -86.6929
5.1919 | -86.6926
5.1922 | -86.6931
5.1925 | -86.6943
5.1929 | -86.6945
5.1932 | -86.6952
5.1935 | -86.6946
5.1938 | -86.6949
5.1941 | -86.6949
5.1944 | -86.6958
5.1947 | -86.6975
5.195 | -86.6989
5.1953 | -86.7006
5.1956 | -86.7017
5.1959 | -86.7022
5.1963 | -86.7024
5.1966 | -86.7027
5.1969 | -86.7029
5.1972 | -86.7033
5.1975 | -86.7046
5.1978 | -86.7054
5.1981 | -86.7059
5.1984 | -86.7075
5.1987 | -86.7078
5.199 | -86.7083
5.1993 | -86.7093
5.1996 | -86.7106
5.2 | -86.7115
5.2003 | -86.7117
5.2006 | -86.7117
5.2009 | -86.7117
5.2012 | -86.7118
5.2015 | -86.7122
5.2018 | -86.7124
5.2021 | -86.7119
5.2024 | -86.7114
5.2027 | -86.7121
5.203 | -86.7132
5.2034 | -86.7133
5.2037 | -86.7132
5.204 | -86.7133
5.2043 | -86.7127
5.2046 | -86.7123
5.2049 | -86.7123
5.2052 | -86.7125
5.2055 | -86.7129
5.2058 | -86.7129
5.2061 | -86.7123
5.2064 | -86.7121
5.2068 | -86.7119
5.2071 | -86.7121
5.2074 | -86.7117
5.2077 | -86.7115
5.208 | -86.7106
5.2083 | -86.7094
5.2086 | -86.7089
5.2089 | -86.7091
5.2092 | -86.7098
5.2095 | -86.71
5.2098 | -86.7103
5.2101 | -86.7107
5.2105 | -86.7102
5.2108 | -86.7092
5.2111 | -86.7094
5.2114 | -86.7088
5.2117 | -86.7095
5.212 | -86.7094
5.2123 | -86.7093
5.2126 | -86.709
5.2129 | -86.7087
5.2132 | -86.7087
5.2135 | -86.7087
5.2139 | -86.7088
5.2142 | -86.7083
5.2145 | -86.7079
5.2148 | -86.7084
5.2151 | -86.7083
5.2154 | -86.7079
5.2157 | -86.7088
5.216 | -86.7099
5.2163 | -86.7106
5.2166 | -86.7116
5.2169 | -86.7116
5.2172 | -86.7122
5.2176 | -86.7129
5.2179 | -86.7132
5.2182 | -86.7135
5.2185 | -86.7138
5.2188 | -86.713
5.2191 | -86.712
5.2194 | -86.7128
5.2197 | -86.7131
5.22 | -86.713
5.2203 | -86.7133
5.2206 | -86.7133
5.221 | -86.7135
5.2213 | -86.7135
5.2216 | -86.7136
5.2219 | -86.7136
5.2222 | -86.7136
5.2225 | -86.7137
5.2228 | -86.7132
5.2231 | -86.7135
5.2234 | -86.7134
5.2237 | -86.7142
5.224 | -86.7153
5.2244 | -86.7154
5.2247 | -86.7151
5.225 | -86.7152
5.2253 | -86.716
5.2256 | -86.7163
5.2259 | -86.7166
5.2262 | -86.7166
5.2265 | -86.7164
5.2268 | -86.7166
5.2271 | -86.7171
5.2274 | -86.7178
5.2277 | -86.7174
5.2281 | -86.7173
5.2284 | -86.7171
5.2287 | -86.7164
5.229 | -86.7159
5.2293 | -86.7161
5.2296 | -86.7161
5.2299 | -86.7161
5.2302 | -86.7183
5.2305 | -86.7226
5.2308 | -86.7248
5.2311 | -86.7251
5.2315 | -86.7253
5.2318 | -86.7252
5.2321 | -86.7249
5.2324 | -86.7246
5.2327 | -86.7241
5.233 | -86.724
5.2333 | -86.7241
5.2336 | -86.724
5.2339 | -86.7241
5.2342 | -86.7242
5.2345 | -86.724
5.2348 | -86.7237
5.2352 | -86.7235
5.2355 | -86.7232
5.2358 | -86.7228
5.2361 | -86.7229
5.2364 | -86.7227
5.2367 | -86.723
5.237 | -86.723
5.2373 | -86.7229
5.2376 | -86.7236
5.2379 | -86.7237
5.2382 | -86.724
5.2386 | -86.7234
5.2389 | -86.7234
5.2392 | -86.7249
5.2395 | -86.7249
5.2398 | -86.7248
5.2401 | -86.7238
5.2404 | -86.7232
5.2407 | -86.7242
5.241 | -86.7255
5.2413 | -86.725
5.2416 | -86.7253
5.242 | -86.7254
5.2423 | -86.7252
5.2426 | -86.7251
5.2429 | -86.7255
5.2432 | -86.7252
5.2435 | -86.7253
5.2438 | -86.725
5.2441 | -86.7241
5.2444 | -86.7244
5.2447 | -86.7239
5.245 | -86.7234
5.2453 | -86.7235
5.2457 | -86.7239
5.246 | -86.7233
5.2463 | -86.7222
5.2466 | -86.7219
5.2469 | -86.7222
5.2472 | -86.7222
5.2475 | -86.7222
5.2478 | -86.7224
5.2481 | -86.7222
5.2484 | -86.722
5.2487 | -86.722
5.2491 | -86.7224
5.2494 | -86.7219
5.2497 | -86.7208
5.25 | -86.7209
5.2503 | -86.721
5.2506 | -86.7209
5.2509 | -86.7193
5.2512 | -86.7174
5.2515 | -86.7168
5.2518 | -86.7165
5.2521 | -86.7149
5.2524 | -86.7124
5.2528 | -86.7121
5.2531 | -86.7131
5.2534 | -86.7138
5.2537 | -86.7139
5.254 | -86.7141
5.2543 | -86.7144
5.2546 | -86.7139
5.2549 | -86.7126
5.2552 | -86.7121
5.2555 | -86.7114
5.2558 | -86.7117
5.2562 | -86.7118
5.2565 | -86.712
5.2568 | -86.7115
5.2571 | -86.7102
5.2574 | -86.7104
5.2577 | -86.7108
5.258 | -86.7108
5.2583 | -86.7106
5.2586 | -86.711
5.2589 | -86.7108
5.2592 | -86.7104
5.2595 | -86.7105
5.2599 | -86.7106
5.2602 | -86.7105
5.2605 | -86.7102
5.2608 | -86.7094
5.2611 | -86.7091
5.2614 | -86.709
5.2617 | -86.7088
5.262 | -86.7087
5.2623 | -86.7081
5.2626 | -86.7074
5.2629 | -86.706
5.2633 | -86.7047
5.2636 | -86.705
5.2639 | -86.7044
5.2642 | -86.7042
5.2645 | -86.7055
5.2648 | -86.7049
5.2651 | -86.7037
5.2654 | -86.7027
5.2657 | -86.7019
5.266 | -86.7013
5.2663 | -86.7009
5.2667 | -86.7
5.267 | -86.6994
5.2673 | -86.6992
5.2676 | -86.6991
5.2679 | -86.6995
5.2682 | -86.6994
5.2685 | -86.6991
5.2688 | -86.6988
5.2691 | -86.6979
5.2694 | -86.6977
5.2697 | -86.6975
5.27 | -86.6973
5.2704 | -86.6971
5.2707 | -86.6962
5.271 | -86.6958
5.2713 | -86.6953
5.2716 | -86.6959
5.2719 | -86.6962
5.2722 | -86.696
5.2725 | -86.6956
5.2728 | -86.6948
5.2731 | -86.6949
5.2734 | -86.6959
5.2738 | -86.6964
5.2741 | -86.6962
5.2744 | -86.6961
5.2747 | -86.6954
5.275 | -86.6951
5.2753 | -86.6952
5.2756 | -86.6945
5.2759 | -86.6941
5.2762 | -86.6935
5.2765 | -86.6928
5.2768 | -86.6924
5.2771 | -86.6921
5.2775 | -86.6923
5.2778 | -86.6929
5.2781 | -86.6937
5.2784 | -86.6942
5.2787 | -86.6941
5.279 | -86.6942
5.2793 | -86.6942
5.2796 | -86.6946
5.2799 | -86.694
5.2802 | -86.6937
5.2805 | -86.6952
5.2809 | -86.6958
5.2812 | -86.6955
5.2815 | -86.6951
5.2818 | -86.6952
5.2821 | -86.6953
5.2824 | -86.6946
5.2827 | -86.6929
5.283 | -86.6929
5.2833 | -86.693
5.2836 | -86.6941
5.2839 | -86.6948
5.2843 | -86.6948
5.2846 | -86.6941
5.2849 | -86.6939
5.2852 | -86.6939
5.2855 | -86.6955
5.2858 | -86.6968
5.2861 | -86.6974
5.2864 | -86.6975
5.2867 | -86.6973
5.287 | -86.6972
5.2873 | -86.6989
5.2876 | -86.6999
5.288 | -86.7002
5.2883 | -86.7005
5.2886 | -86.7008
5.2889 | -86.701
5.2892 | -86.7009
5.2895 | -86.7012
5.2898 | -86.7016
5.2901 | -86.7014
5.2904 | -86.7014
5.2907 | -86.7013
5.291 | -86.7009
5.2914 | -86.7003
5.2917 | -86.7003
5.292 | -86.6992
5.2923 | -86.6983
5.2926 | -86.6996
5.2929 | -86.7015
5.2932 | -86.702
5.2935 | -86.7018
5.2938 | -86.7015
5.2941 | -86.7005
5.2944 | -86.7001
5.2947 | -86.7
5.2951 | -86.7001
5.2954 | -86.6998
5.2957 | -86.7
5.296 | -86.7001
5.2963 | -86.701
5.2966 | -86.7006
5.2969 | -86.6997
5.2972 | -86.6993
5.2975 | -86.6986
5.2978 | -86.6988
5.2981 | -86.6999
5.2985 | -86.7003
5.2988 | -86.7009
5.2991 | -86.7008
5.2994 | -86.7011
5.2997 | -86.7018
5.3 | -86.7023
5.3003 | -86.702
5.3006 | -86.7027
5.3009 | -86.7028
5.3012 | -86.7033
5.3015 | -86.7035
5.3019 | -86.7034
5.3022 | -86.7033
5.3025 | -86.7037
5.3028 | -86.7036
5.3031 | -86.7038
5.3034 | -86.7039
5.3037 | -86.7039
5.304 | -86.7041
5.3043 | -86.704
5.3046 | -86.7043
5.3049 | -86.7039
5.3052 | -86.7037
5.3056 | -86.7037
5.3059 | -86.7039
5.3062 | -86.7042
5.3065 | -86.7052
5.3068 | -86.7057
5.3071 | -86.7058
5.3074 | -86.7058
5.3077 | -86.7059
5.308 | -86.7061
5.3083 | -86.7062
5.3086 | -86.7061
5.309 | -86.7052
5.3093 | -86.7044
5.3096 | -86.7049
5.3099 | -86.7033
5.3102 | -86.704
5.3105 | -86.7051
5.3108 | -86.7053
5.3111 | -86.7051
5.3114 | -86.7053
5.3117 | -86.7058
5.312 | -86.7065
5.3123 | -86.7058
5.3127 | -86.7047
5.313 | -86.7041
5.3133 | -86.7036
5.3136 | -86.704
5.3139 | -86.704
5.3142 | -86.7038
5.3145 | -86.7038
5.3148 | -86.7038
5.3151 | -86.7039
5.3154 | -86.7038
5.3157 | -86.7038
5.3161 | -86.7028
5.3164 | -86.7021
5.3167 | -86.7022
5.317 | -86.7024
5.3173 | -86.7026
5.3176 | -86.7033
5.3179 | -86.7039
5.3182 | -86.7041
5.3185 | -86.7041
5.3188 | -86.7043
5.3191 | -86.7045
5.3195 | -86.7045
5.3198 | -86.7049
5.3201 | -86.7051
5.3204 | -86.7056
5.3207 | -86.7063
5.321 | -86.7085
5.3213 | -86.7099
5.3216 | -86.7105
5.3219 | -86.7112
5.3222 | -86.7115
5.3225 | -86.7118
5.3228 | -86.7116
5.3232 | -86.7118
5.3235 | -86.7122
5.3238 | -86.7131
5.3241 | -86.7133
5.3244 | -86.7132
5.3247 | -86.7123
5.325 | -86.7115
5.3253 | -86.7115
5.3256 | -86.7117
5.3259 | -86.7123
5.3262 | -86.7126
5.3266 | -86.7128
5.3269 | -86.7131
5.3272 | -86.7142
5.3275 | -86.7146
5.3278 | -86.7154
5.3281 | -86.7162
5.3284 | -86.7165
5.3287 | -86.7162
5.329 | -86.7165
5.3293 | -86.7165
5.3296 | -86.7167
5.3299 | -86.7168
5.3303 | -86.7168
5.3306 | -86.7168
5.3309 | -86.7174
5.3312 | -86.7181
5.3315 | -86.7178
5.3318 | -86.7169
5.3321 | -86.7169
5.3324 | -86.717
5.3327 | -86.717
5.333 | -86.7176
5.3333 | -86.7175
5.3337 | -86.7188
5.334 | -86.7197
5.3343 | -86.7197
5.3346 | -86.72
5.3349 | -86.72
5.3352 | -86.7201
5.3355 | -86.7202
5.3358 | -86.72
5.3361 | -86.7209
5.3364 | -86.7216
5.3367 | -86.7223
5.3371 | -86.7234
5.3374 | -86.7237
5.3377 | -86.724
5.338 | -86.7243
5.3383 | -86.7246
5.3386 | -86.7248
5.3389 | -86.7252
5.3392 | -86.7258
5.3395 | -86.7262
5.3398 | -86.7276
5.3401 | -86.7289
5.3404 | -86.7288
5.3408 | -86.729
5.3411 | -86.7289
5.3414 | -86.7281
5.3417 | -86.7279
5.342 | -86.7282
5.3423 | -86.7286
5.3426 | -86.7284
5.3429 | -86.7278
5.3432 | -86.7274
5.3435 | -86.7271
5.3438 | -86.729
5.3442 | -86.7299
5.3445 | -86.7298
5.3448 | -86.7302
5.3451 | -86.7308
5.3454 | -86.7313
5.3457 | -86.7317
5.346 | -86.7322
5.3463 | -86.7325
5.3466 | -86.7334
5.3469 | -86.7341
5.3472 | -86.7343
5.3475 | -86.7348
5.3479 | -86.7352
5.3482 | -86.7357
5.3485 | -86.7364
5.3488 | -86.7367
5.3491 | -86.7372
5.3494 | -86.7378
5.3497 | -86.7382
5.35 | -86.7386
5.3503 | -86.7375
5.3506 | -86.7369
5.3509 | -86.7375
5.3513 | -86.7378
5.3516 | -86.7379
5.3519 | -86.7383
5.3522 | -86.7383
5.3525 | -86.7382
5.3528 | -86.7383
5.3531 | -86.7389
5.3534 | -86.7393
5.3537 | -86.7401
5.354 | -86.7395
5.3543 | -86.7376
5.3547 | -86.7374
5.355 | -86.7381
5.3553 | -86.7381
5.3556 | -86.738
5.3559 | -86.7378
5.3562 | -86.7384
5.3565 | -86.7402
5.3568 | -86.7424
5.3571 | -86.7428
5.3574 | -86.7429
5.3577 | -86.7434
5.358 | -86.7439
5.3584 | -86.7448
5.3587 | -86.745
5.359 | -86.746
5.3593 | -86.7468
5.3596 | -86.7463
5.3599 | -86.7466
5.3602 | -86.748
5.3605 | -86.7499
5.3608 | -86.7503
5.3611 | -86.7506
5.3614 | -86.7505
5.3618 | -86.7505
5.3621 | -86.7509
5.3624 | -86.7527
5.3627 | -86.7539
5.363 | -86.7549
5.3633 | -86.755
5.3636 | -86.7552
5.3639 | -86.7562
5.3642 | -86.7561
5.3645 | -86.7559
5.3648 | -86.7563
5.3651 | -86.7568
5.3655 | -86.7575
5.3658 | -86.7595
5.3661 | -86.7616
5.3664 | -86.7627
5.3667 | -86.7632
5.367 | -86.7638
5.3673 | -86.7646
5.3676 | -86.7646
5.3679 | -86.7647
5.3682 | -86.7651
5.3685 | -86.7649
5.3689 | -86.7653
5.3692 | -86.7659
5.3695 | -86.7666
5.3698 | -86.7672
5.3701 | -86.7671
5.3704 | -86.7666
5.3707 | -86.7674
5.371 | -86.7678
5.3713 | -86.7681
5.3716 | -86.7687
5.3719 | -86.7694
5.3723 | -86.7702
5.3726 | -86.7684
5.3729 | -86.7674
5.3732 | -86.767
5.3735 | -86.7673
5.3738 | -86.7681
5.3741 | -86.769
5.3744 | -86.7696
5.3747 | -86.77
5.375 | -86.7703
5.3753 | -86.7709
5.3756 | -86.7717
5.376 | -86.7723
5.3763 | -86.7731
5.3766 | -86.7742
5.3769 | -86.7756
5.3772 | -86.7761
5.3775 | -86.7756
5.3778 | -86.7751
5.3781 | -86.7753
5.3784 | -86.7759
5.3787 | -86.776
5.379 | -86.7757
5.3794 | -86.7768
5.3797 | -86.7769
5.38 | -86.7771
5.3803 | -86.7782
5.3806 | -86.7792
5.3809 | -86.781
5.3812 | -86.7818
5.3815 | -86.7823
5.3818 | -86.7832
5.3821 | -86.7836
5.3824 | -86.7851
5.3827 | -86.7863
5.3831 | -86.787
5.3834 | -86.7874
5.3837 | -86.787
5.384 | -86.7869
5.3843 | -86.7873
5.3846 | -86.7883
5.3849 | -86.791
5.3852 | -86.7915
5.3855 | -86.7904
5.3858 | -86.7904
5.3861 | -86.7908
5.3865 | -86.7913
5.3868 | -86.7916
5.3871 | -86.7916
5.3874 | -86.7921
5.3877 | -86.7926
5.388 | -86.7928
5.3883 | -86.7933
5.3886 | -86.7948
5.3889 | -86.7963
5.3892 | -86.796
5.3895 | -86.7945
5.3899 | -86.7954
5.3902 | -86.7963
5.3905 | -86.7965
5.3908 | -86.7958
5.3911 | -86.7962
5.3914 | -86.7968
5.3917 | -86.7969
5.392 | -86.7971
5.3923 | -86.7974
5.3926 | -86.798
5.3929 | -86.7993
5.3932 | -86.8008
5.3936 | -86.8018
5.3939 | -86.8022
5.3942 | -86.8035
5.3945 | -86.8047
5.3948 | -86.8052
5.3951 | -86.8058
5.3954 | -86.8065
5.3957 | -86.8067
5.396 | -86.8069
5.3963 | -86.8079
5.3966 | -86.8095
5.397 | -86.8096
5.3973 | -86.8096
5.3976 | -86.8095
5.3979 | -86.8099
5.3982 | -86.81
5.3985 | -86.811
5.3988 | -86.8116
5.3991 | -86.8121
5.3994 | -86.8124
5.3997 | -86.8121
5.4 | -86.8125
5.4003 | -86.8117
5.4007 | -86.81
5.401 | -86.8091
5.4013 | -86.8083
5.4016 | -86.808
5.4019 | -86.808
5.4022 | -86.8084
5.4025 | -86.8087
5.4028 | -86.8088
5.4031 | -86.809
5.4034 | -86.8093
5.4037 | -86.8097
5.4041 | -86.8101
5.4044 | -86.8105
5.4047 | -86.8109
5.405 | -86.8112
5.4053 | -86.8116
5.4056 | -86.8123
5.4059 | -86.8133
5.4062 | -86.8141
5.4065 | -86.8141
5.4068 | -86.8139
5.4071 | -86.8147
5.4074 | -86.8156
5.4078 | -86.8161
5.4081 | -86.8162
5.4084 | -86.8168
5.4087 | -86.8175
5.409 | -86.818
5.4093 | -86.8186
5.4096 | -86.8195
5.4099 | -86.8208
5.4102 | -86.8209
5.4105 | -86.8209
5.4108 | -86.8213
5.4112 | -86.822
5.4115 | -86.8216
5.4118 | -86.8205
5.4121 | -86.8208
5.4124 | -86.8213
5.4127 | -86.8216
5.413 | -86.8219
5.4133 | -86.8224
5.4136 | -86.8228
5.4139 | -86.8233
5.4142 | -86.824
5.4146 | -86.8241
5.4149 | -86.8243
5.4152 | -86.8254
5.4155 | -86.8264
5.4158 | -86.8272
5.4161 | -86.8275
5.4164 | -86.8284
5.4167 | -86.8289
5.417 | -86.8297
5.4173 | -86.8305
5.4176 | -86.8303
5.4179 | -86.8304
5.4183 | -86.8309
5.4186 | -86.8313
5.4189 | -86.8316
5.4192 | -86.8318
5.4195 | -86.8324
5.4198 | -86.8337
5.4201 | -86.8349
5.4204 | -86.8356
5.4207 | -86.8357
5.421 | -86.836
5.4213 | -86.8362
5.4217 | -86.8364
5.422 | -86.836
5.4223 | -86.836
5.4226 | -86.8368
5.4229 | -86.8374
5.4232 | -86.838
5.4235 | -86.8388
5.4238 | -86.8394
5.4241 | -86.8398
5.4244 | -86.8406
5.4247 | -86.8412
5.425 | -86.8409
5.4254 | -86.8404
5.4257 | -86.8406
5.426 | -86.8413
5.4263 | -86.8429
5.4266 | -86.8443
5.4269 | -86.8443
5.4272 | -86.8446
5.4275 | -86.8447
5.4278 | -86.8453
5.4281 | -86.8467
5.4284 | -86.8469
5.4288 | -86.8479
5.4291 | -86.8482
5.4294 | -86.8479
5.4297 | -86.8485
5.43 | -86.8488
5.4303 | -86.8491
5.4306 | -86.8491
5.4309 | -86.8492
5.4312 | -86.8496
5.4315 | -86.8501
5.4318 | -86.8504
5.4322 | -86.85
5.4325 | -86.8508
5.4328 | -86.8513
5.4331 | -86.8518
5.4334 | -86.8533
5.4337 | -86.8539
5.434 | -86.8542
5.4343 | -86.8541
5.4346 | -86.8537
5.4349 | -86.8541
5.4352 | -86.8544
5.4355 | -86.8556
5.4359 | -86.856
5.4362 | -86.8558
5.4365 | -86.8557
5.4368 | -86.8558
5.4371 | -86.8562
5.4374 | -86.8564
5.4377 | -86.8555
5.438 | -86.8555
5.4383 | -86.8557
5.4386 | -86.8561
5.4389 | -86.8567
5.4393 | -86.8569
5.4396 | -86.8571
5.4399 | -86.8577
5.4402 | -86.8587
5.4405 | -86.8591
5.4408 | -86.8595
5.4411 | -86.8596
5.4414 | -86.8603
5.4417 | -86.861
5.442 | -86.8608
5.4423 | -86.86
5.4426 | -86.8603
5.443 | -86.86
5.4433 | -86.863
5.4436 | -86.8652
5.4439 | -86.8652
5.4442 | -86.8652
5.4445 | -86.8654
5.4448 | -86.8652
5.4451 | -86.8653
5.4454 | -86.8654
5.4457 | -86.8663
5.446 | -86.8674
5.4464 | -86.8673
5.4467 | -86.8672
5.447 | -86.8673
5.4473 | -86.868
5.4476 | -86.8687
5.4479 | -86.8685
5.4482 | -86.8686
5.4485 | -86.8677
5.4488 | -86.8686
5.4491 | -86.8692
5.4494 | -86.8689
5.4498 | -86.869
5.4501 | -86.8698
5.4504 | -86.8718
5.4507 | -86.8751
5.451 | -86.8772
5.4513 | -86.8773
5.4516 | -86.8776
5.4519 | -86.8774
5.4522 | -86.8779
5.4525 | -86.8784
5.4528 | -86.8782
5.4531 | -86.8783
5.4535 | -86.8783
5.4538 | -86.8777
5.4541 | -86.8774
5.4544 | -86.8775
5.4547 | -86.8781
5.455 | -86.8789
5.4553 | -86.8788
5.4556 | -86.8793
5.4559 | -86.8793
5.4562 | -86.8796
5.4565 | -86.88
5.4569 | -86.8804
5.4572 | -86.8803
5.4575 | -86.8806
5.4578 | -86.8812
5.4581 | -86.8816
5.4584 | -86.8814
5.4587 | -86.8816
5.459 | -86.8811
5.4593 | -86.8808
5.4596 | -86.8811
5.4599 | -86.8808
5.4602 | -86.8801
5.4606 | -86.8794
5.4609 | -86.8787
5.4612 | -86.8784
5.4615 | -86.8784
5.4618 | -86.8786
5.4621 | -86.8791
5.4624 | -86.8791
5.4627 | -86.8798
5.463 | -86.88
5.4633 | -86.8798
5.4636 | -86.8802
5.464 | -86.8812
5.4643 | -86.8816
5.4646 | -86.8819
5.4649 | -86.8821
5.4652 | -86.8822
5.4655 | -86.8824
5.4658 | -86.8822
5.4661 | -86.8815
5.4664 | -86.8811
5.4667 | -86.8807
5.467 | -86.8797
5.4674 | -86.8784
5.4677 | -86.8784
5.468 | -86.8785
5.4683 | -86.8788
5.4686 | -86.8793
5.4689 | -86.8796
5.4692 | -86.8801
5.4695 | -86.8808
5.4698 | -86.8801
5.4701 | -86.8808
5.4704 | -86.8812
5.4707 | -86.8821
5.4711 | -86.8827
5.4714 | -86.8838
5.4717 | -86.8839
5.472 | -86.8834
5.4723 | -86.8835
5.4726 | -86.8841
5.4729 | -86.8843
5.4732 | -86.8843
5.4735 | -86.8838
5.4738 | -86.8836
5.4741 | -86.8839
5.4745 | -86.8834
5.4748 | -86.8844
5.4751 | -86.8859
5.4754 | -86.8864
5.4757 | -86.8862
5.476 | -86.8856
5.4763 | -86.8854
5.4766 | -86.885
5.4769 | -86.8848
5.4772 | -86.8854
5.4775 | -86.8866
5.4778 | -86.8867
5.4782 | -86.8867
5.4785 | -86.8868
5.4788 | -86.8876
5.4791 | -86.8877
5.4794 | -86.8876
5.4797 | -86.8878
5.48 | -86.8883
5.4803 | -86.8879
5.4806 | -86.8877
5.4809 | -86.8877
5.4812 | -86.8875
5.4816 | -86.8875
5.4819 | -86.8874
5.4822 | -86.8876
5.4825 | -86.8877
5.4828 | -86.8881
5.4831 | -86.8893
5.4834 | -86.8902
5.4837 | -86.8904
5.484 | -86.8904
5.4843 | -86.8906
5.4846 | -86.8899
5.485 | -86.8893
5.4853 | -86.889
5.4856 | -86.8892
5.4859 | -86.8888
5.4862 | -86.8887
5.4865 | -86.8891
5.4868 | -86.8893
5.4871 | -86.8892
5.4874 | -86.8901
5.4877 | -86.8894
5.488 | -86.8892
5.4883 | -86.8893
5.4887 | -86.8894
5.489 | -86.8895
5.4893 | -86.8896
5.4896 | -86.8894
5.4899 | -86.8888
5.4902 | -86.8878
5.4905 | -86.8874
5.4908 | -86.8873
5.4911 | -86.8875
5.4914 | -86.8882
5.4917 | -86.8889
5.4921 | -86.8894
5.4924 | -86.8896
5.4927 | -86.8897
5.493 | -86.8899
5.4933 | -86.8903
5.4936 | -86.891
5.4939 | -86.891
5.4942 | -86.8905
5.4945 | -86.8907
5.4948 | -86.8908
5.4951 | -86.891
5.4954 | -86.8911
5.4958 | -86.892
5.4961 | -86.894
5.4964 | -86.894
5.4967 | -86.8955
5.497 | -86.8962
5.4973 | -86.8959
5.4976 | -86.8956
5.4979 | -86.8963
5.4982 | -86.8966
5.4985 | -86.8963
5.4988 | -86.8962
5.4992 | -86.8961
5.4995 | -86.8962
5.4998 | -86.8967
5.5001 | -86.8973
5.5004 | -86.8979
5.5007 | -86.8984
5.501 | -86.899
5.5013 | -86.8991
5.5016 | -86.8991
5.5019 | -86.8993
5.5022 | -86.8996
5.5026 | -86.8992
5.5029 | -86.8993
5.5032 | -86.8998
5.5035 | -86.8997
5.5038 | -86.8999
5.5041 | -86.9001
5.5044 | -86.8994
5.5047 | -86.8976
5.505 | -86.8972
5.5053 | -86.8978
5.5056 | -86.8979
5.5059 | -86.8977
5.5063 | -86.8966
5.5066 | -86.895
5.5069 | -86.8943
5.5072 | -86.8938
5.5075 | -86.894
5.5078 | -86.8942
5.5081 | -86.8946
5.5084 | -86.8949
5.5087 | -86.895
5.509 | -86.8956
5.5093 | -86.8977
5.5097 | -86.899
5.51 | -86.8996
5.5103 | -86.8997
5.5106 | -86.9004
5.5109 | -86.9006
5.5112 | -86.9006
5.5115 | -86.9009
5.5118 | -86.9017
5.5121 | -86.9029
5.5124 | -86.9032
5.5127 | -86.9035
5.513 | -86.903
5.5134 | -86.9036
5.5137 | -86.9041
5.514 | -86.9043
5.5143 | -86.9049
5.5146 | -86.9051
5.5149 | -86.9059
5.5152 | -86.9064
5.5155 | -86.9062
5.5158 | -86.9054
5.5161 | -86.9054
5.5164 | -86.9056
5.5168 | -86.9056
5.5171 | -86.9053
5.5174 | -86.9056
5.5177 | -86.9059
5.518 | -86.9062
5.5183 | -86.9063
5.5186 | -86.9063
5.5189 | -86.9064
5.5192 | -86.9065
5.5195 | -86.9066
5.5198 | -86.9074
5.5202 | -86.9086
5.5205 | -86.9083
5.5208 | -86.908
5.5211 | -86.9093
5.5214 | -86.9133
5.5217 | -86.9148
5.522 | -86.9149
5.5223 | -86.9146
5.5226 | -86.9151
5.5229 | -86.9151
5.5232 | -86.9154
5.5235 | -86.9157
5.5239 | -86.9158
5.5242 | -86.9152
5.5245 | -86.9151
5.5248 | -86.9155
5.5251 | -86.9156
5.5254 | -86.9151
5.5257 | -86.9152
5.526 | -86.9151
5.5263 | -86.9152
5.5266 | -86.9154
5.5269 | -86.9162
5.5273 | -86.9169
5.5276 | -86.9175
5.5279 | -86.9178
5.5282 | -86.9179
5.5285 | -86.9187
5.5288 | -86.9207
5.5291 | -86.9207
5.5294 | -86.9201
5.5297 | -86.9201
5.53 | -86.9203
5.5303 | -86.9212
5.5306 | -86.9235
5.531 | -86.9255
5.5313 | -86.9267
5.5316 | -86.927
5.5319 | -86.927
5.5322 | -86.9272
5.5325 | -86.9278
5.5328 | -86.9274
5.5331 | -86.927
5.5334 | -86.9273
5.5337 | -86.9274
5.534 | -86.9283
5.5344 | -86.9302
5.5347 | -86.9306
5.535 | -86.93
5.5353 | -86.9291
5.5356 | -86.9282
5.5359 | -86.928
5.5362 | -86.9301
5.5365 | -86.9317
5.5368 | -86.9332
5.5371 | -86.9346
5.5374 | -86.9351
5.5377 | -86.9366
5.5381 | -86.9374
5.5384 | -86.9375
5.5387 | -86.9379
5.539 | -86.9387
5.5393 | -86.939
5.5396 | -86.9394
5.5399 | -86.9404
5.5402 | -86.941
5.5405 | -86.9414
5.5408 | -86.9417
5.5411 | -86.9429
5.5415 | -86.9436
5.5418 | -86.9437
5.5421 | -86.9428
5.5424 | -86.9417
5.5427 | -86.9409
5.543 | -86.9405
5.5433 | -86.9408
5.5436 | -86.9408
5.5439 | -86.9406
5.5442 | -86.9415
5.5445 | -86.9427
5.5449 | -86.9434
5.5452 | -86.9437
5.5455 | -86.9444
5.5458 | -86.9445
5.5461 | -86.9434
5.5464 | -86.9433
5.5467 | -86.9436
5.547 | -86.9434
5.5473 | -86.9428
5.5476 | -86.9419
5.5479 | -86.9418
5.5482 | -86.9424
5.5486 | -86.9433
5.5489 | -86.9444
5.5492 | -86.9445
5.5495 | -86.9447
5.5498 | -86.9455
5.5501 | -86.9468
5.5504 | -86.9471
5.5507 | -86.9464
5.551 | -86.9466
5.5513 | -86.947
5.5516 | -86.9476
5.552 | -86.9479
5.5523 | -86.9483
5.5526 | -86.9486
5.5529 | -86.9488
5.5532 | -86.9498
5.5535 | -86.9507
5.5538 | -86.9511
5.5541 | -86.9515
5.5544 | -86.9523
5.5547 | -86.9531
5.555 | -86.9538
5.5553 | -86.9553
5.5557 | -86.9557
5.556 | -86.9564
5.5563 | -86.9568
5.5566 | -86.9575
5.5569 | -86.9579
5.5572 | -86.9582
5.5575 | -86.9583
5.5578 | -86.9588
5.5581 | -86.9582
5.5584 | -86.958
5.5587 | -86.9593
5.5591 | -86.9602
5.5594 | -86.9606
5.5597 | -86.9612
5.56 | -86.9614
5.5603 | -86.962
5.5606 | -86.9631
5.5609 | -86.9644
5.5612 | -86.966
5.5615 | -86.9663
5.5618 | -86.9661
5.5621 | -86.9665
5.5625 | -86.9671
5.5628 | -86.968
5.5631 | -86.9682
5.5634 | -86.9677
5.5637 | -86.9679
5.564 | -86.9687
5.5643 | -86.969
5.5646 | -86.97
5.5649 | -86.9701
5.5652 | -86.971
5.5655 | -86.9716
5.5658 | -86.9716
5.5662 | -86.9724
5.5665 | -86.973
5.5668 | -86.9738
5.5671 | -86.9733
5.5674 | -86.9738
5.5677 | -86.9742
5.568 | -86.9753
5.5683 | -86.9761
5.5686 | -86.9764
5.5689 | -86.9767
5.5692 | -86.9773
5.5696 | -86.9776
5.5699 | -86.9783
5.5702 | -86.9788
5.5705 | -86.9791
5.5708 | -86.979
5.5711 | -86.9788
5.5714 | -86.9785
5.5717 | -86.9794
5.572 | -86.9803
5.5723 | -86.9811
5.5726 | -86.9819
5.5729 | -86.9824
5.5733 | -86.9829
5.5736 | -86.9835
5.5739 | -86.9843
5.5742 | -86.985
5.5745 | -86.9851
5.5748 | -86.9852
5.5751 | -86.9856
5.5754 | -86.9859
5.5757 | -86.9865
5.576 | -86.9877
5.5763 | -86.9879
5.5767 | -86.9879
5.577 | -86.9888
5.5773 | -86.9892
5.5776 | -86.9899
5.5779 | -86.9907
5.5782 | -86.9917
5.5785 | -86.9927
5.5788 | -86.9938
5.5791 | -86.9946
5.5794 | -86.9955
5.5797 | -86.9959
5.5801 | -86.9953
5.5804 | -86.9949
5.5807 | -86.9953
5.581 | -86.996
5.5813 | -86.9965
5.5816 | -86.997
5.5819 | -86.9976
5.5822 | -86.9981
5.5825 | -86.9996
5.5828 | -87.0023
5.5831 | -87.0038
5.5834 | -87.0041
5.5838 | -87.0045
5.5841 | -87.005
5.5844 | -87.0055
5.5847 | -87.0058
5.585 | -87.0061
5.5853 | -87.0067
5.5856 | -87.0075
5.5859 | -87.0078
5.5862 | -87.0074
5.5865 | -87.0085
5.5868 | -87.0095
5.5872 | -87.0102
5.5875 | -87.011
5.5878 | -87.0118
5.5881 | -87.0113
5.5884 | -87.0101
5.5887 | -87.0105
5.589 | -87.0112
5.5893 | -87.0121
5.5896 | -87.0127
5.5899 | -87.0133
5.5902 | -87.0131
5.5905 | -87.0128
5.5909 | -87.0126
5.5912 | -87.0126
5.5915 | -87.0128
5.5918 | -87.0124
5.5921 | -87.0106
5.5924 | -87.0097
5.5927 | -87.0103
5.593 | -87.0107
5.5933 | -87.0101
5.5936 | -87.01
5.5939 | -87.0112
5.5943 | -87.0124
5.5946 | -87.0126
5.5949 | -87.0129
5.5952 | -87.0141
5.5955 | -87.0152
5.5958 | -87.0157
5.5961 | -87.0158
5.5964 | -87.0147
5.5967 | -87.014
5.597 | -87.0148
5.5973 | -87.0139
5.5977 | -87.0125
5.598 | -87.012
5.5983 | -87.0122
5.5986 | -87.013
5.5989 | -87.0129
5.5992 | -87.0129
5.5995 | -87.0134
5.5998 | -87.014
5.6001 | -87.0144
5.6004 | -87.0142
5.6007 | -87.0146
5.601 | -87.0151
5.6014 | -87.0144
5.6017 | -87.0125
5.602 | -87.0123
5.6023 | -87.0127
5.6026 | -87.0128
5.6029 | -87.0125
5.6032 | -87.013
5.6035 | -87.0137
5.6038 | -87.015
5.6041 | -87.0154
5.6044 | -87.0146
5.6048 | -87.0141
5.6051 | -87.0143
5.6054 | -87.0141
5.6057 | -87.0136
5.606 | -87.0135
5.6063 | -87.0129
5.6066 | -87.0128
5.6069 | -87.0129
5.6072 | -87.0129
5.6075 | -87.0121
5.6078 | -87.0108
5.6081 | -87.0107
5.6085 | -87.0099
5.6088 | -87.01
5.6091 | -87.01
5.6094 | -87.0105
5.6097 | -87.0108
5.61 | -87.0109
5.6103 | -87.0113
5.6106 | -87.0115
5.6109 | -87.0117
5.6112 | -87.0123
5.6115 | -87.0121
5.6119 | -87.0123
5.6122 | -87.0135
5.6125 | -87.0137
5.6128 | -87.0137
5.6131 | -87.0139
5.6134 | -87.0139
5.6137 | -87.0137
5.614 | -87.0138
5.6143 | -87.014
5.6146 | -87.0141
5.6149 | -87.0141
5.6153 | -87.0142
5.6156 | -87.015
5.6159 | -87.0151
5.6162 | -87.0147
5.6165 | -87.0143
5.6168 | -87.0142
5.6171 | -87.0147
5.6174 | -87.0146
5.6177 | -87.0146
5.618 | -87.0152
5.6183 | -87.0157
5.6186 | -87.0157
5.619 | -87.016
5.6193 | -87.0164
5.6196 | -87.0166
5.6199 | -87.0169
5.6202 | -87.0166
5.6205 | -87.0149
5.6208 | -87.0141
5.6211 | -87.0142
5.6214 | -87.0151
5.6217 | -87.0136
5.622 | -87.0137
5.6224 | -87.0142
5.6227 | -87.0142
5.623 | -87.0138
5.6233 | -87.0133
5.6236 | -87.0139
5.6239 | -87.0141
5.6242 | -87.0142
5.6245 | -87.0145
5.6248 | -87.0147
5.6251 | -87.0143
5.6254 | -87.0137
5.6257 | -87.0145
5.6261 | -87.0147
5.6264 | -87.0142
5.6267 | -87.0137
5.627 | -87.0142
5.6273 | -87.0143
5.6276 | -87.0135
5.6279 | -87.0122
5.6282 | -87.0116
5.6285 | -87.0134
5.6288 | -87.0149
5.6291 | -87.0155
5.6295 | -87.0151
5.6298 | -87.0164
5.6301 | -87.0165
5.6304 | -87.0186
5.6307 | -87.0189
5.631 | -87.0186
5.6313 | -87.0183
5.6316 | -87.0178
5.6319 | -87.0171
5.6322 | -87.0167
5.6325 | -87.017
5.6329 | -87.0167
5.6332 | -87.0158
5.6335 | -87.0162
5.6338 | -87.0155
5.6341 | -87.0142
5.6344 | -87.0136
5.6347 | -87.0131
5.635 | -87.0133
5.6353 | -87.0138
5.6356 | -87.0135
5.6359 | -87.0138
5.6362 | -87.0135
5.6366 | -87.0134
5.6369 | -87.0131
5.6372 | -87.013
5.6375 | -87.0131
5.6378 | -87.013
5.6381 | -87.0128
5.6384 | -87.0122
5.6387 | -87.0113
5.639 | -87.0104
5.6393 | -87.0107
5.6396 | -87.0098
5.64 | -87.0093
5.6403 | -87.0095
5.6406 | -87.0092
5.6409 | -87.009
5.6412 | -87.0087
5.6415 | -87.0095
5.6418 | -87.0102
5.6421 | -87.0094
5.6424 | -87.0087
5.6427 | -87.008
5.643 | -87.0075
5.6433 | -87.0061
5.6437 | -87.0055
5.644 | -87.0063
5.6443 | -87.0064
5.6446 | -87.0066
5.6449 | -87.0064
5.6452 | -87.0058
5.6455 | -87.0043
5.6458 | -87.0032
5.6461 | -87.0028
5.6464 | -87.0023
5.6467 | -87.0019
5.6471 | -87.0016
5.6474 | -87.0011
5.6477 | -87.0006
5.648 | -87
5.6483 | -87.0002
5.6486 | -86.9989
5.6489 | -86.9979
5.6492 | -86.9967
5.6495 | -86.9964
5.6498 | -86.9969
5.6501 | -86.9973
5.6505 | -86.997
5.6508 | -86.996
5.6511 | -86.9955
5.6514 | -86.9948
5.6517 | -86.9947
5.652 | -86.9945
5.6523 | -86.9937
5.6526 | -86.9932
5.6529 | -86.9925
5.6532 | -86.9909
5.6535 | -86.9878
5.6538 | -86.9855
5.6542 | -86.9849
5.6545 | -86.984
5.6548 | -86.9827
5.6551 | -86.9822
5.6554 | -86.982
5.6557 | -86.9812
5.656 | -86.9806
5.6563 | -86.9797
5.6566 | -86.979
5.6569 | -86.9773
5.6572 | -86.9758
5.6576 | -86.9747
5.6579 | -86.9739
5.6582 | -86.9731
5.6585 | -86.9723
5.6588 | -86.9713
5.6591 | -86.9708
5.6594 | -86.9691
5.6597 | -86.968
5.66 | -86.9674
5.6603 | -86.9665
5.6606 | -86.9652
5.6609 | -86.9644
5.6613 | -86.9637
5.6616 | -86.9628
5.6619 | -86.962
5.6622 | -86.9623
5.6625 | -86.9629
5.6628 | -86.9627
5.6631 | -86.9616
5.6634 | -86.9603
5.6637 | -86.9594
5.664 | -86.9584
5.6643 | -86.9575
5.6647 | -86.9567
5.665 | -86.955
5.6653 | -86.9539
5.6656 | -86.9528
5.6659 | -86.9505
5.6662 | -86.9493
5.6665 | -86.9484
5.6668 | -86.9473
5.6671 | -86.9458
5.6674 | -86.9448
5.6677 | -86.9442
5.6681 | -86.9435
5.6684 | -86.9429
5.6687 | -86.942
5.669 | -86.9412
5.6693 | -86.9405
5.6696 | -86.9394
5.6699 | -86.9392
5.6702 | -86.9385
5.6705 | -86.9376
5.6708 | -86.9366
5.6711 | -86.9352
5.6714 | -86.9332
5.6718 | -86.9307
5.6721 | -86.9299
5.6724 | -86.9291
5.6727 | -86.9283
5.673 | -86.9279
5.6733 | -86.9269
5.6736 | -86.9256
5.6739 | -86.9242
5.6742 | -86.9211
5.6745 | -86.9198
5.6748 | -86.9181
5.6752 | -86.9179
5.6755 | -86.9171
5.6758 | -86.9155
5.6761 | -86.9146
5.6764 | -86.914
5.6767 | -86.9126
5.677 | -86.9122
5.6773 | -86.9113
5.6776 | -86.9107
5.6779 | -86.9102
5.6782 | -86.9094
5.6785 | -86.9091
5.6789 | -86.9085
5.6792 | -86.9077
5.6795 | -86.9069
5.6798 | -86.9056
5.6801 | -86.9041
5.6804 | -86.9021
5.6807 | -86.9
5.681 | -86.8984
5.6813 | -86.8976
5.6816 | -86.8962
5.6819 | -86.894
5.6823 | -86.8918
5.6826 | -86.8903
5.6829 | -86.8886
5.6832 | -86.8873
5.6835 | -86.8855
5.6838 | -86.8843
5.6841 | -86.8827
5.6844 | -86.8816
5.6847 | -86.8802
5.685 | -86.8786
5.6853 | -86.8769
5.6856 | -86.8754
5.686 | -86.874
5.6863 | -86.8729
5.6866 | -86.8729
5.6869 | -86.8717
5.6872 | -86.8704
5.6875 | -86.8694
5.6878 | -86.8683
5.6881 | -86.8667
5.6884 | -86.8651
5.6887 | -86.8632
5.689 | -86.8614
5.6894 | -86.8596
5.6897 | -86.8583
5.69 | -86.8569
5.6903 | -86.8557
5.6906 | -86.8552
5.6909 | -86.8544
5.6912 | -86.8527
5.6915 | -86.8515
5.6918 | -86.8512
5.6921 | -86.8495
5.6924 | -86.8471
5.6928 | -86.8459
5.6931 | -86.8446
5.6934 | -86.843
5.6937 | -86.8416
5.694 | -86.8405
5.6943 | -86.8392
5.6946 | -86.8387
5.6949 | -86.8385
5.6952 | -86.8378
5.6955 | -86.837
5.6958 | -86.8359
5.6961 | -86.8346
5.6965 | -86.8334
5.6968 | -86.8317
5.6971 | -86.8301
5.6974 | -86.8284
5.6977 | -86.8272
5.698 | -86.8256
5.6983 | -86.8247
5.6986 | -86.8237
5.6989 | -86.8213
5.6992 | -86.8192
5.6995 | -86.8175
5.6999 | -86.8166
5.7002 | -86.8157
5.7005 | -86.8153
5.7008 | -86.8138
5.7011 | -86.8124
5.7014 | -86.8108
5.7017 | -86.8094
5.702 | -86.8079
5.7023 | -86.8067
5.7026 | -86.8049
5.7029 | -86.8033
5.7032 | -86.8016
5.7036 | -86.7999
5.7039 | -86.7981
5.7042 | -86.7962
5.7045 | -86.7933
5.7048 | -86.7908
5.7051 | -86.7887
5.7054 | -86.7867
5.7057 | -86.7851
5.706 | -86.784
5.7063 | -86.7829
5.7066 | -86.7812
5.707 | -86.7798
5.7073 | -86.7783
5.7076 | -86.777
5.7079 | -86.7752
5.7082 | -86.7736
5.7085 | -86.7724
5.7088 | -86.771
5.7091 | -86.7692
5.7094 | -86.7675
5.7097 | -86.7661
5.71 | -86.7638
5.7104 | -86.7618
5.7107 | -86.7595
5.711 | -86.7578
5.7113 | -86.7561
5.7116 | -86.7549
5.7119 | -86.7533
5.7122 | -86.752
5.7125 | -86.7499
5.7128 | -86.747
5.7131 | -86.7448
5.7134 | -86.7434
5.7137 | -86.7418
5.7141 | -86.7399
5.7144 | -86.7377
5.7147 | -86.736
5.715 | -86.7351
5.7153 | -86.7349
5.7156 | -86.7335
5.7159 | -86.732
5.7162 | -86.7305
5.7165 | -86.729
5.7168 | -86.7277
5.7171 | -86.726
5.7175 | -86.7245
5.7178 | -86.7231
5.7181 | -86.721
5.7184 | -86.7193
5.7187 | -86.7182
5.719 | -86.7168
5.7193 | -86.7154
5.7196 | -86.7134
5.7199 | -86.7117
5.7202 | -86.7102
5.7205 | -86.7087
5.7208 | -86.7071
5.7212 | -86.7054
5.7215 | -86.704
5.7218 | -86.7025
5.7221 | -86.701
5.7224 | -86.6995
5.7227 | -86.6982
5.723 | -86.6965
5.7233 | -86.6948
5.7236 | -86.6931
5.7239 | -86.6913
5.7242 | -86.69
5.7246 | -86.6888
5.7249 | -86.6873
5.7252 | -86.6858
5.7255 | -86.6852
5.7258 | -86.6841
5.7261 | -86.6826
5.7264 | -86.6813
5.7267 | -86.68
5.727 | -86.6793
5.7273 | -86.6786
5.7276 | -86.6781
5.728 | -86.6768
5.7283 | -86.6757
5.7286 | -86.6741
5.7289 | -86.6731
5.7292 | -86.6714
5.7295 | -86.6687
5.7298 | -86.6668
5.7301 | -86.6652
5.7304 | -86.6638
5.7307 | -86.6623
5.731 | -86.6603
5.7313 | -86.6579
5.7317 | -86.6557
5.732 | -86.6539
5.7323 | -86.6524
5.7326 | -86.651
5.7329 | -86.6492
5.7332 | -86.6471
5.7335 | -86.6457
5.7338 | -86.6438
5.7341 | -86.6421
5.7344 | -86.6406
5.7347 | -86.6393
5.7351 | -86.6389
5.7354 | -86.6384
5.7357 | -86.6367
5.736 | -86.6351
5.7363 | -86.6339
5.7366 | -86.6329
5.7369 | -86.632
5.7372 | -86.6309
5.7375 | -86.6293
5.7378 | -86.6267
5.7381 | -86.6249
5.7384 | -86.6235
5.7388 | -86.622
5.7391 | -86.6199
5.7394 | -86.6186
5.7397 | -86.6182
5.74 | -86.6173
5.7403 | -86.6156
5.7406 | -86.6144
5.7409 | -86.6134
5.7412 | -86.6119
5.7415 | -86.6103
5.7418 | -86.6088
5.7422 | -86.6071
5.7425 | -86.6052
5.7428 | -86.604
5.7431 | -86.602
5.7434 | -86.6001
5.7437 | -86.5984
5.744 | -86.5971
5.7443 | -86.5954
5.7446 | -86.5938
5.7449 | -86.5926
5.7452 | -86.591
5.7456 | -86.5895
5.7459 | -86.5887
5.7462 | -86.5871
5.7465 | -86.5861
5.7468 | -86.5851
5.7471 | -86.584
5.7474 | -86.5827
5.7477 | -86.5812
5.748 | -86.5789
5.7483 | -86.5748
5.7486 | -86.5722
5.7489 | -86.5706
5.7493 | -86.5688
5.7496 | -86.5675
5.7499 | -86.5661
5.7502 | -86.5647
5.7505 | -86.5641
5.7508 | -86.5626
5.7511 | -86.561
5.7514 | -86.5596
5.7517 | -86.558
5.752 | -86.5562
5.7523 | -86.5545
5.7527 | -86.5525
5.753 | -86.55
5.7533 | -86.5475
5.7536 | -86.5458
5.7539 | -86.5442
5.7542 | -86.5427
5.7545 | -86.5417
5.7548 | -86.5389
5.7551 | -86.5368
5.7554 | -86.5354
5.7557 | -86.5341
5.756 | -86.5328
5.7564 | -86.5316
5.7567 | -86.5304
5.757 | -86.528
5.7573 | -86.526
5.7576 | -86.5247
5.7579 | -86.5237
5.7582 | -86.5227
5.7585 | -86.5213
5.7588 | -86.5201
5.7591 | -86.5209
5.7594 | -86.521
5.7598 | -86.5199
5.7601 | -86.5187
5.7604 | -86.5172
5.7607 | -86.5159
5.761 | -86.5144
5.7613 | -86.5131
5.7616 | -86.5117
5.7619 | -86.5101
5.7622 | -86.5089
5.7625 | -86.5078
5.7628 | -86.5063
5.7632 | -86.5052
5.7635 | -86.5036
5.7638 | -86.501
5.7641 | -86.4997
5.7644 | -86.4992
5.7647 | -86.4982
5.765 | -86.496
5.7653 | -86.4935
5.7656 | -86.4917
5.7659 | -86.4907
5.7662 | -86.4894
5.7665 | -86.489
5.7669 | -86.4873
5.7672 | -86.4844
5.7675 | -86.4821
5.7678 | -86.4797
5.7681 | -86.4776
5.7684 | -86.476
5.7687 | -86.475
5.769 | -86.4739
5.7693 | -86.4725
5.7696 | -86.4706
5.7699 | -86.4685
5.7703 | -86.4671
5.7706 | -86.4658
5.7709 | -86.465
5.7712 | -86.464
5.7715 | -86.4618
5.7718 | -86.4598
5.7721 | -86.458
5.7724 | -86.4569
5.7727 | -86.4563
5.773 | -86.4563
5.7733 | -86.455
5.7736 | -86.4536
5.774 | -86.452
5.7743 | -86.4498
5.7746 | -86.4486
5.7749 | -86.4473
5.7752 | -86.4466
5.7755 | -86.4458
5.7758 | -86.4444
5.7761 | -86.442
5.7764 | -86.4393
5.7767 | -86.4382
5.777 | -86.4378
5.7774 | -86.437
5.7777 | -86.4358
5.778 | -86.4346
5.7783 | -86.4336
5.7786 | -86.4335
5.7789 | -86.4323
5.7792 | -86.4312
5.7795 | -86.429
5.7798 | -86.4269
5.7801 | -86.4259
5.7804 | -86.4249
5.7808 | -86.4238
5.7811 | -86.4225
5.7814 | -86.4213
5.7817 | -86.4211
5.782 | -86.42
5.7823 | -86.419
5.7826 | -86.4179
5.7829 | -86.4172
5.7832 | -86.4163
5.7835 | -86.4152
5.7838 | -86.4145
5.7841 | -86.4129
5.7845 | -86.4115
5.7848 | -86.4098
5.7851 | -86.4083
5.7854 | -86.4075
5.7857 | -86.4068
5.786 | -86.4051
5.7863 | -86.4043
5.7866 | -86.4033
5.7869 | -86.4022
5.7872 | -86.4012
5.7875 | -86.4
5.7879 | -86.3988
5.7882 | -86.3983
5.7885 | -86.397
5.7888 | -86.3965
5.7891 | -86.3963
5.7894 | -86.3962
5.7897 | -86.3959
5.79 | -86.3949
5.7903 | -86.3938
5.7906 | -86.394
5.7909 | -86.3926
5.7912 | -86.3916
5.7916 | -86.3908
5.7919 | -86.3893
5.7922 | -86.3877
5.7925 | -86.3867
5.7928 | -86.3857
5.7931 | -86.3849
5.7934 | -86.384
5.7937 | -86.3832
5.794 | -86.382
5.7943 | -86.3805
5.7946 | -86.3794
5.795 | -86.3785
5.7953 | -86.3784
5.7956 | -86.3782
5.7959 | -86.3771
5.7962 | -86.376
5.7965 | -86.375
5.7968 | -86.374
5.7971 | -86.3733
5.7974 | -86.3724
5.7977 | -86.3713
5.798 | -86.3699
5.7984 | -86.369
5.7987 | -86.3678
5.799 | -86.3666
5.7993 | -86.3664
5.7996 | -86.3656
5.7999 | -86.3646
5.8002 | -86.3639
5.8005 | -86.3628
5.8008 | -86.3627
5.8011 | -86.3628
5.8014 | -86.3619
5.8017 | -86.3609
5.8021 | -86.3599
5.8024 | -86.3591
5.8027 | -86.3579
5.803 | -86.3565
5.8033 | -86.3552
5.8036 | -86.3542
5.8039 | -86.353
5.8042 | -86.3522
5.8045 | -86.3514
5.8048 | -86.3507
5.8051 | -86.3501
5.8055 | -86.3497
5.8058 | -86.3491
5.8061 | -86.3479
5.8064 | -86.3469
5.8067 | -86.3462
5.807 | -86.3451
5.8073 | -86.3448
5.8076 | -86.3444
5.8079 | -86.3438
5.8082 | -86.3439
5.8085 | -86.3432
5.8088 | -86.3428
5.8092 | -86.3424
5.8095 | -86.3419
5.8098 | -86.3407
5.8101 | -86.3396
5.8104 | -86.3393
5.8107 | -86.3393
5.811 | -86.3386
5.8113 | -86.338
5.8116 | -86.3375
5.8119 | -86.3375
5.8122 | -86.3372
5.8126 | -86.3366
5.8129 | -86.3361
5.8132 | -86.3356
5.8135 | -86.3352
5.8138 | -86.3348
5.8141 | -86.3341
5.8144 | -86.3331
5.8147 | -86.3327
5.815 | -86.3331
5.8153 | -86.3333
5.8156 | -86.3326
5.8159 | -86.3309
5.8163 | -86.3302
5.8166 | -86.3297
5.8169 | -86.329
5.8172 | -86.329
5.8175 | -86.3284
5.8178 | -86.3283
5.8181 | -86.3284
5.8184 | -86.328
5.8187 | -86.327
5.819 | -86.3268
5.8193 | -86.3262
5.8197 | -86.3253
5.82 | -86.3254
5.8203 | -86.3247
5.8206 | -86.3241
5.8209 | -86.3234
5.8212 | -86.3233
5.8215 | -86.3229
5.8218 | -86.3225
5.8221 | -86.3221
5.8224 | -86.3222
5.8227 | -86.3229
5.8231 | -86.3228
5.8234 | -86.3226
5.8237 | -86.3225
5.824 | -86.3222
5.8243 | -86.3209
5.8246 | -86.3203
5.8249 | -86.3201
5.8252 | -86.3201
5.8255 | -86.3205
5.8258 | -86.3202
5.8261 | -86.3208
5.8264 | -86.3207
5.8268 | -86.3204
5.8271 | -86.3202
5.8274 | -86.3204
5.8277 | -86.3197
5.828 | -86.3186
5.8283 | -86.318
5.8286 | -86.3174
5.8289 | -86.3183
5.8292 | -86.3185
5.8295 | -86.3183
5.8298 | -86.3183
5.8302 | -86.3182
5.8305 | -86.3185
5.8308 | -86.3181
5.8311 | -86.3181
5.8314 | -86.3177
5.8317 | -86.318
5.832 | -86.3182
5.8323 | -86.3183
5.8326 | -86.3183
5.8329 | -86.3181
5.8332 | -86.3185
5.8335 | -86.3192
5.8339 | -86.3198
5.8342 | -86.3204
5.8345 | -86.3206
5.8348 | -86.3202
5.8351 | -86.3215
5.8354 | -86.3215
5.8357 | -86.3208
5.836 | -86.3194
5.8363 | -86.3192
5.8366 | -86.3196
5.8369 | -86.3196
5.8373 | -86.3197
5.8376 | -86.3197
5.8379 | -86.3197
5.8382 | -86.3201
5.8385 | -86.3204
5.8388 | -86.3205
5.8391 | -86.3203
5.8394 | -86.3205
5.8397 | -86.3206
5.84 | -86.3207
5.8403 | -86.3221
5.8407 | -86.3226
5.841 | -86.3229
5.8413 | -86.323
5.8416 | -86.3223
5.8419 | -86.3224
5.8422 | -86.3225
5.8425 | -86.3226
5.8428 | -86.3224
5.8431 | -86.3221
5.8434 | -86.3218
5.8437 | -86.321
5.844 | -86.3213
5.8444 | -86.3212
5.8447 | -86.3212
5.845 | -86.3216
5.8453 | -86.3221
5.8456 | -86.3223
5.8459 | -86.3227
5.8462 | -86.3234
5.8465 | -86.3239
5.8468 | -86.3243
5.8471 | -86.3243
5.8474 | -86.3244
5.8478 | -86.3245
5.8481 | -86.3243
5.8484 | -86.3247
5.8487 | -86.3254
5.849 | -86.3255
5.8493 | -86.3254
5.8496 | -86.3257
5.8499 | -86.3259
5.8502 | -86.3262
5.8505 | -86.3269
5.8508 | -86.3279
5.8511 | -86.3283
5.8515 | -86.3291
5.8518 | -86.3304
5.8521 | -86.3313
5.8524 | -86.3316
5.8527 | -86.3314
5.853 | -86.3325
5.8533 | -86.3331
5.8536 | -86.3336
5.8539 | -86.3343
5.8542 | -86.3346
5.8545 | -86.3342
5.8549 | -86.3345
5.8552 | -86.3341
5.8555 | -86.3339
5.8558 | -86.3342
5.8561 | -86.3348
5.8564 | -86.335
5.8567 | -86.3347
5.857 | -86.3348
5.8573 | -86.335
5.8576 | -86.335
5.8579 | -86.3353
5.8583 | -86.3356
5.8586 | -86.3354
5.8589 | -86.3347
5.8592 | -86.3354
5.8595 | -86.3362
5.8598 | -86.337
5.8601 | -86.3386
5.8604 | -86.3402
5.8607 | -86.3408
5.861 | -86.341
5.8613 | -86.341
5.8616 | -86.3422
5.862 | -86.3424
5.8623 | -86.3423
5.8626 | -86.3427
5.8629 | -86.3432
5.8632 | -86.3445
5.8635 | -86.3453
5.8638 | -86.3453
5.8641 | -86.3454
5.8644 | -86.3459
5.8647 | -86.3467
5.865 | -86.347
5.8654 | -86.3475
5.8657 | -86.3484
5.866 | -86.3488
5.8663 | -86.3489
5.8666 | -86.3495
5.8669 | -86.3498
5.8672 | -86.3504
5.8675 | -86.3507
5.8678 | -86.3512
5.8681 | -86.352
5.8684 | -86.3524
5.8687 | -86.353
5.8691 | -86.3535
5.8694 | -86.3536
5.8697 | -86.3543
5.87 | -86.3553
5.8703 | -86.3562
5.8706 | -86.3578
5.8709 | -86.3583
5.8712 | -86.3586
5.8715 | -86.3589
5.8718 | -86.3596
5.8721 | -86.3603
5.8725 | -86.361
5.8728 | -86.3608
5.8731 | -86.3604
5.8734 | -86.3601
5.8737 | -86.3609
5.874 | -86.3616
5.8743 | -86.3624
5.8746 | -86.3633
5.8749 | -86.3642
5.8752 | -86.365
5.8755 | -86.3661
5.8759 | -86.3661
5.8762 | -86.3664
5.8765 | -86.3671
5.8768 | -86.3678
5.8771 | -86.3685
5.8774 | -86.3696
5.8777 | -86.3702
5.878 | -86.3709
5.8783 | -86.3716
5.8786 | -86.3724
5.8789 | -86.3738
5.8792 | -86.3753
5.8796 | -86.3767
5.8799 | -86.3776
5.8802 | -86.3783
5.8805 | -86.3794
5.8808 | -86.3805
5.8811 | -86.3814
5.8814 | -86.3823
5.8817 | -86.3829
5.882 | -86.383
5.8823 | -86.3839
5.8826 | -86.3854
5.883 | -86.3864
5.8833 | -86.3869
5.8836 | -86.3879
5.8839 | -86.3887
5.8842 | -86.3895
5.8845 | -86.3904
5.8848 | -86.3914
5.8851 | -86.393
5.8854 | -86.394
5.8857 | -86.3945
5.886 | -86.3955
5.8863 | -86.3964
5.8867 | -86.3971
5.887 | -86.3978
5.8873 | -86.398
5.8876 | -86.399
5.8879 | -86.3993
5.8882 | -86.4003
5.8885 | -86.4014
5.8888 | -86.4024
5.8891 | -86.4033
5.8894 | -86.4046
5.8897 | -86.4057
5.8901 | -86.4067
5.8904 | -86.4073
5.8907 | -86.4074
5.891 | -86.4075
5.8913 | -86.4087
5.8916 | -86.4097
5.8919 | -86.4109
5.8922 | -86.4121
5.8925 | -86.4125
5.8928 | -86.4134
5.8931 | -86.4149
5.8934 | -86.4159
5.8938 | -86.4168
5.8941 | -86.4181
5.8944 | -86.4191
5.8947 | -86.4201
5.895 | -86.421
5.8953 | -86.4218
5.8956 | -86.4229
5.8959 | -86.4239
5.8962 | -86.4247
5.8965 | -86.4252
5.8968 | -86.4264
5.8972 | -86.4269
5.8975 | -86.427
5.8978 | -86.4277
5.8981 | -86.4283
5.8984 | -86.4296
5.8987 | -86.4305
5.899 | -86.4306
5.8993 | -86.4306
5.8996 | -86.4313
5.8999 | -86.432
5.9002 | -86.4336
5.9006 | -86.4362
5.9009 | -86.4377
5.9012 | -86.4386
5.9015 | -86.4393
5.9018 | -86.4396
5.9021 | -86.4406
5.9024 | -86.4427
5.9027 | -86.4455
5.903 | -86.4479
5.9033 | -86.4498
5.9036 | -86.451
5.9039 | -86.452
5.9043 | -86.4526
5.9046 | -86.4532
5.9049 | -86.4538
5.9052 | -86.4544
5.9055 | -86.4554
5.9058 | -86.4563
5.9061 | -86.4571
5.9064 | -86.4587
5.9067 | -86.4596
5.907 | -86.4596
5.9073 | -86.4597
5.9077 | -86.4601
5.908 | -86.4602
5.9083 | -86.4609
5.9086 | -86.4619
5.9089 | -86.4624
5.9092 | -86.463
5.9095 | -86.4649
5.9098 | -86.4662
5.9101 | -86.4671
5.9104 | -86.4677
5.9107 | -86.4681
5.9111 | -86.4692
5.9114 | -86.4695
5.9117 | -86.4697
5.912 | -86.4704
5.9123 | -86.4716
5.9126 | -86.4732
5.9129 | -86.4741
5.9132 | -86.4743
5.9135 | -86.4747
5.9138 | -86.476
5.9141 | -86.4769
5.9144 | -86.4779
5.9148 | -86.479
5.9151 | -86.4803
5.9154 | -86.4812
5.9157 | -86.4827
5.916 | -86.4847
5.9163 | -86.4863
5.9166 | -86.4872
5.9169 | -86.4881
5.9172 | -86.4891
5.9175 | -86.4906
5.9178 | -86.492
5.9182 | -86.4931
5.9185 | -86.4941
5.9188 | -86.4943
5.9191 | -86.4955
5.9194 | -86.4971
5.9197 | -86.4981
5.92 | -86.4994
5.9203 | -86.5
5.9206 | -86.5003
5.9209 | -86.5013
5.9212 | -86.502
5.9215 | -86.5035
5.9219 | -86.5054
5.9222 | -86.5071
5.9225 | -86.5083
5.9228 | -86.5098
5.9231 | -86.51
5.9234 | -86.5107
5.9237 | -86.5119
5.924 | -86.5128
5.9243 | -86.513
5.9246 | -86.5136
5.9249 | -86.5144
5.9253 | -86.5145
5.9256 | -86.5147
5.9259 | -86.5157
5.9262 | -86.5166
5.9265 | -86.5175
5.9268 | -86.5179
5.9271 | -86.5191
5.9274 | -86.5203
5.9277 | -86.5214
5.928 | -86.5226
5.9283 | -86.5237
5.9287 | -86.5245
5.929 | -86.5254
5.9293 | -86.5267
5.9296 | -86.5279
5.9299 | -86.5294
5.9302 | -86.5307
5.9305 | -86.5316
5.9308 | -86.5333
5.9311 | -86.5348
5.9314 | -86.5359
5.9317 | -86.537
5.932 | -86.5382
5.9324 | -86.5395
5.9327 | -86.5408
5.933 | -86.5416
5.9333 | -86.5437
5.9336 | -86.545
5.9339 | -86.5468
5.9342 | -86.5486
5.9345 | -86.5498
5.9348 | -86.5508
5.9351 | -86.5515
5.9354 | -86.5518
5.9358 | -86.5527
5.9361 | -86.5537
5.9364 | -86.5548
5.9367 | -86.5563
5.937 | -86.5578
5.9373 | -86.5585
5.9376 | -86.5596
5.9379 | -86.5605
5.9382 | -86.5615
5.9385 | -86.5623
5.9388 | -86.5632
5.9391 | -86.5642
5.9395 | -86.5654
5.9398 | -86.5666
5.9401 | -86.5678
5.9404 | -86.5692
5.9407 | -86.5714
5.941 | -86.5737
5.9413 | -86.5759
5.9416 | -86.5775
5.9419 | -86.5792
5.9422 | -86.5801
5.9425 | -86.5816
5.9429 | -86.5835
5.9432 | -86.5848
5.9435 | -86.5855
5.9438 | -86.5865
5.9441 | -86.5875
5.9444 | -86.5886
5.9447 | -86.5899
5.945 | -86.5922
5.9453 | -86.5937
5.9456 | -86.5947
5.9459 | -86.5963
5.9463 | -86.5983
5.9466 | -86.6011
5.9469 | -86.6033
5.9472 | -86.6046
5.9475 | -86.6057
5.9478 | -86.6069
5.9481 | -86.6083
5.9484 | -86.6096
5.9487 | -86.6108
5.949 | -86.6119
5.9493 | -86.6127
5.9496 | -86.6137
5.95 | -86.616
5.9503 | -86.6176
5.9506 | -86.6186
5.9509 | -86.6202
5.9512 | -86.6212
5.9515 | -86.6223
5.9518 | -86.6232
5.9521 | -86.6239
5.9524 | -86.6247
5.9527 | -86.6255
5.953 | -86.6258
5.9534 | -86.6269
5.9537 | -86.6281
5.954 | -86.6295
5.9543 | -86.6308
5.9546 | -86.6319
5.9549 | -86.6328
5.9552 | -86.6338
5.9555 | -86.6352
5.9558 | -86.6366
5.9561 | -86.6379
5.9564 | -86.639
5.9567 | -86.6397
5.9571 | -86.6411
5.9574 | -86.6422
5.9577 | -86.6441
5.958 | -86.646
5.9583 | -86.647
5.9586 | -86.6483
5.9589 | -86.6497
5.9592 | -86.6511
5.9595 | -86.6522
5.9598 | -86.6532
5.9601 | -86.6548
5.9605 | -86.6562
5.9608 | -86.6575
5.9611 | -86.6587
5.9614 | -86.6592
5.9617 | -86.6599
5.962 | -86.6616
5.9623 | -86.6624
5.9626 | -86.664
5.9629 | -86.665
5.9632 | -86.6665
5.9635 | -86.6684
5.9638 | -86.6692
5.9642 | -86.6698
5.9645 | -86.6707
5.9648 | -86.6719
5.9651 | -86.6738
5.9654 | -86.6751
5.9657 | -86.6768
5.966 | -86.6787
5.9663 | -86.6798
5.9666 | -86.6812
5.9669 | -86.6829
5.9672 | -86.685
5.9676 | -86.6863
5.9679 | -86.6877
5.9682 | -86.6888
5.9685 | -86.6899
5.9688 | -86.6918
5.9691 | -86.693
5.9694 | -86.6945
5.9697 | -86.6957
5.97 | -86.697
5.9703 | -86.6982
5.9706 | -86.699
5.971 | -86.7002
5.9713 | -86.7017
5.9716 | -86.7037
5.9719 | -86.7056
5.9722 | -86.7077
5.9725 | -86.7086
5.9728 | -86.7091
5.9731 | -86.7102
5.9734 | -86.7116
5.9737 | -86.7122
5.974 | -86.7126
5.9743 | -86.7138
5.9747 | -86.7148
5.975 | -86.7157
5.9753 | -86.7174
5.9756 | -86.719
5.9759 | -86.7206
5.9762 | -86.7221
5.9765 | -86.7236
5.9768 | -86.7246
5.9771 | -86.7268
5.9774 | -86.7285
5.9777 | -86.7306
5.9781 | -86.7309
5.9784 | -86.7314
5.9787 | -86.7327
5.979 | -86.7339
5.9793 | -86.7352
5.9796 | -86.7362
5.9799 | -86.7374
5.9802 | -86.7393
5.9805 | -86.7408
5.9808 | -86.7412
5.9811 | -86.7422
5.9814 | -86.7446
5.9818 | -86.7471
5.9821 | -86.7493
5.9824 | -86.7505
5.9827 | -86.7522
5.983 | -86.7536
5.9833 | -86.7553
5.9836 | -86.757
5.9839 | -86.7586
5.9842 | -86.7597
5.9845 | -86.7611
5.9848 | -86.7626
5.9852 | -86.764
5.9855 | -86.7653
5.9858 | -86.7666
5.9861 | -86.7682
5.9864 | -86.77
5.9867 | -86.7716
5.987 | -86.7725
5.9873 | -86.7732
5.9876 | -86.7746
5.9879 | -86.7766
5.9882 | -86.7791
5.9886 | -86.7801
5.9889 | -86.7809
5.9892 | -86.782
5.9895 | -86.7832
5.9898 | -86.7845
5.9901 | -86.787
5.9904 | -86.7895
5.9907 | -86.7908
5.991 | -86.7921
5.9913 | -86.7936
5.9916 | -86.7953
5.9919 | -86.796
5.9923 | -86.7975
5.9926 | -86.7992
5.9929 | -86.8006
5.9932 | -86.8024
5.9935 | -86.804
5.9938 | -86.8059
5.9941 | -86.8093
5.9944 | -86.8117
5.9947 | -86.8127
5.995 | -86.8139
5.9953 | -86.8156
5.9957 | -86.8177
5.996 | -86.8194
5.9963 | -86.821
5.9966 | -86.823
5.9969 | -86.8245
5.9972 | -86.8259
5.9975 | -86.8273
5.9978 | -86.8286
5.9981 | -86.8298
5.9984 | -86.8312
5.9987 | -86.8327
5.999 | -86.8352
5.9994 | -86.8372
5.9997 | -86.8381
6 | -86.8406
6.0003 | -86.8433
6.0006 | -86.845
6.0009 | -86.8465
6.0012 | -86.8475
6.0015 | -86.8486
6.0018 | -86.8493
6.0021 | -86.8505
6.0024 | -86.8516
6.0028 | -86.8529
6.0031 | -86.8548
6.0034 | -86.8564
6.0037 | -86.8577
6.004 | -86.8599
6.0043 | -86.8614
6.0046 | -86.8646
6.0049 | -86.866
6.0052 | -86.8671
6.0055 | -86.8687
6.0058 | -86.8706
6.0062 | -86.8719
6.0065 | -86.8735
6.0068 | -86.8745
6.0071 | -86.8755
6.0074 | -86.877
6.0077 | -86.8797
6.008 | -86.8821
6.0083 | -86.884
6.0086 | -86.8859
6.0089 | -86.8876
6.0092 | -86.889
6.0095 | -86.8914
6.0099 | -86.893
6.0102 | -86.8949
6.0105 | -86.8965
6.0108 | -86.8988
6.0111 | -86.9016
6.0114 | -86.9033
6.0117 | -86.906
6.012 | -86.908
6.0123 | -86.9093
6.0126 | -86.9104
6.0129 | -86.9118
6.0133 | -86.9136
6.0136 | -86.9148
6.0139 | -86.9161
6.0142 | -86.9175
6.0145 | -86.9186
6.0148 | -86.92
6.0151 | -86.9221
6.0154 | -86.9234
6.0157 | -86.9243
6.016 | -86.9256
6.0163 | -86.9258
6.0166 | -86.9262
6.017 | -86.9268
6.0173 | -86.9283
6.0176 | -86.9302
6.0179 | -86.9316
6.0182 | -86.9328
6.0185 | -86.9343
6.0188 | -86.9356
6.0191 | -86.9373
6.0194 | -86.9387
6.0197 | -86.94
6.02 | -86.9413
6.0204 | -86.9433
6.0207 | -86.9449
6.021 | -86.9468
6.0213 | -86.9485
6.0216 | -86.9503
6.0219 | -86.9516
6.0222 | -86.9536
6.0225 | -86.9547
6.0228 | -86.9557
6.0231 | -86.9571
6.0234 | -86.9585
6.0237 | -86.9598
6.0241 | -86.9613
6.0244 | -86.9634
6.0247 | -86.9664
6.025 | -86.9687
6.0253 | -86.9703
6.0256 | -86.9711
6.0259 | -86.9727
6.0262 | -86.9751
6.0265 | -86.9761
6.0268 | -86.9771
6.0271 | -86.9783
6.0275 | -86.9791
6.0278 | -86.9809
6.0281 | -86.9827
6.0284 | -86.9845
6.0287 | -86.9861
6.029 | -86.9873
6.0293 | -86.9887
6.0296 | -86.9905
6.0299 | -86.9923
6.0302 | -86.9936
6.0305 | -86.9952
6.0309 | -86.997
6.0312 | -86.9987
6.0315 | -87.0002
6.0318 | -87.0017
6.0321 | -87.0036
6.0324 | -87.0052
6.0327 | -87.0084
6.033 | -87.011
6.0333 | -87.0115
6.0336 | -87.012
6.0339 | -87.0128
6.0342 | -87.0146
6.0346 | -87.0152
6.0349 | -87.0155
6.0352 | -87.0167
6.0355 | -87.018
6.0358 | -87.0187
6.0361 | -87.0202
6.0364 | -87.0217
6.0367 | -87.0232
6.037 | -87.0251
6.0373 | -87.0267
6.0376 | -87.0278
6.038 | -87.0302
6.0383 | -87.0327
6.0386 | -87.034
6.0389 | -87.036
6.0392 | -87.0385
6.0395 | -87.0403
6.0398 | -87.0424
6.0401 | -87.0432
6.0404 | -87.0434
6.0407 | -87.0446
6.041 | -87.0461
6.0414 | -87.047
6.0417 | -87.0485
6.042 | -87.05
6.0423 | -87.0514
6.0426 | -87.0513
6.0429 | -87.0506
6.0432 | -87.0516
6.0435 | -87.0532
6.0438 | -87.055
6.0441 | -87.0566
6.0444 | -87.0569
6.0447 | -87.0573
6.0451 | -87.0593
6.0454 | -87.0619
6.0457 | -87.0639
6.046 | -87.0654
6.0463 | -87.0663
6.0466 | -87.0674
6.0469 | -87.0686
6.0472 | -87.0696
6.0475 | -87.0708
6.0478 | -87.0728
6.0481 | -87.0749
6.0485 | -87.0768
6.0488 | -87.0787
6.0491 | -87.0799
6.0494 | -87.082
6.0497 | -87.0834
6.05 | -87.084
6.0503 | -87.0853
6.0506 | -87.0862
6.0509 | -87.0875
6.0512 | -87.0893
6.0515 | -87.092
6.0518 | -87.0942
6.0522 | -87.095
6.0525 | -87.0963
6.0528 | -87.0978
6.0531 | -87.099
6.0534 | -87.1004
6.0537 | -87.102
6.054 | -87.1035
6.0543 | -87.1048
6.0546 | -87.1058
6.0549 | -87.1077
6.0552 | -87.1096
6.0556 | -87.1116
6.0559 | -87.1128
6.0562 | -87.1141
6.0565 | -87.1154
6.0568 | -87.1173
6.0571 | -87.1192
6.0574 | -87.1203
6.0577 | -87.122
6.058 | -87.1236
6.0583 | -87.1251
6.0586 | -87.1266
6.059 | -87.128
6.0593 | -87.1288
6.0596 | -87.1299
6.0599 | -87.1319
6.0602 | -87.1359
6.0605 | -87.1387
6.0608 | -87.1403
6.0611 | -87.1419
6.0614 | -87.1437
6.0617 | -87.1447
6.062 | -87.1457
6.0623 | -87.1472
6.0627 | -87.149
6.063 | -87.1508
6.0633 | -87.1523
6.0636 | -87.1542
6.0639 | -87.1559
6.0642 | -87.1579
6.0645 | -87.1597
6.0648 | -87.1613
6.0651 | -87.1629
6.0654 | -87.1645
6.0657 | -87.1661
6.0661 | -87.1675
6.0664 | -87.1697
6.0667 | -87.1701
6.067 | -87.1717
6.0673 | -87.1731
6.0676 | -87.1745
6.0679 | -87.1753
6.0682 | -87.1756
6.0685 | -87.1762
6.0688 | -87.1774
6.0691 | -87.1789
6.0694 | -87.1793
6.0698 | -87.1797
6.0701 | -87.1809
6.0704 | -87.1825
6.0707 | -87.1843
6.071 | -87.1853
6.0713 | -87.1865
6.0716 | -87.188
6.0719 | -87.1898
6.0722 | -87.1919
6.0725 | -87.1934
6.0728 | -87.1937
6.0732 | -87.1943
6.0735 | -87.1958
6.0738 | -87.198
6.0741 | -87.1992
6.0744 | -87.2008
6.0747 | -87.2026
6.075 | -87.2036
6.0753 | -87.2043
6.0756 | -87.2046
6.0759 | -87.2046
6.0762 | -87.2056
6.0766 | -87.2059
6.0769 | -87.2047
6.0772 | -87.205
6.0775 | -87.2064
6.0778 | -87.2075
6.0781 | -87.2083
6.0784 | -87.2091
6.0787 | -87.2105
6.079 | -87.2122
6.0793 | -87.2147
6.0796 | -87.2181
6.0799 | -87.2191
6.0803 | -87.2208
6.0806 | -87.2223
6.0809 | -87.2233
6.0812 | -87.2246
6.0815 | -87.2259
6.0818 | -87.2268
6.0821 | -87.2296
6.0824 | -87.2313
6.0827 | -87.2322
6.083 | -87.234
6.0833 | -87.2353
6.0837 | -87.2363
6.084 | -87.2372
6.0843 | -87.2381
6.0846 | -87.2394
6.0849 | -87.2402
6.0852 | -87.2417
6.0855 | -87.243
6.0858 | -87.2441
6.0861 | -87.245
6.0864 | -87.2462
6.0867 | -87.2473
6.087 | -87.2483
6.0874 | -87.2494
6.0877 | -87.2503
6.088 | -87.2515
6.0883 | -87.2531
6.0886 | -87.2546
6.0889 | -87.2567
6.0892 | -87.2588
6.0895 | -87.2604
6.0898 | -87.2612
6.0901 | -87.2617
6.0904 | -87.2637
6.0908 | -87.2659
6.0911 | -87.2671
6.0914 | -87.2688
6.0917 | -87.2701
6.092 | -87.271
6.0923 | -87.2716
6.0926 | -87.2728
6.0929 | -87.2739
6.0932 | -87.2751
6.0935 | -87.276
6.0938 | -87.277
6.0941 | -87.2783
6.0945 | -87.2799
6.0948 | -87.2813
6.0951 | -87.2825
6.0954 | -87.2837
6.0957 | -87.2847
6.096 | -87.2861
6.0963 | -87.2873
6.0966 | -87.2884
6.0969 | -87.2896
6.0972 | -87.2911
6.0975 | -87.2924
6.0979 | -87.294
6.0982 | -87.2958
6.0985 | -87.2972
6.0988 | -87.2972
6.0991 | -87.2981
6.0994 | -87.2994
6.0997 | -87.3008
6.1 | -87.302
6.1003 | -87.3037
6.1006 | -87.3051
6.1009 | -87.3063
6.1013 | -87.3076
6.1016 | -87.309
6.1019 | -87.3107
6.1022 | -87.3115
6.1025 | -87.3127
6.1028 | -87.3137
6.1031 | -87.315
6.1034 | -87.3163
6.1037 | -87.3174
6.104 | -87.318
6.1043 | -87.3169
6.1046 | -87.3149
6.105 | -87.315
6.1053 | -87.316
6.1056 | -87.3171
6.1059 | -87.3178
6.1062 | -87.3184
6.1065 | -87.32
6.1068 | -87.3215
6.1071 | -87.3228
6.1074 | -87.3239
6.1077 | -87.3249
6.108 | -87.3254
6.1084 | -87.326
6.1087 | -87.3277
6.109 | -87.329
6.1093 | -87.3303
6.1096 | -87.3311
6.1099 | -87.3316
6.1102 | -87.3323
6.1105 | -87.3332
6.1108 | -87.334
6.1111 | -87.3352
6.1114 | -87.3365
6.1117 | -87.3376
6.1121 | -87.3385
6.1124 | -87.3396
6.1127 | -87.3409
6.113 | -87.3429
6.1133 | -87.3445
6.1136 | -87.3454
6.1139 | -87.3463
6.1142 | -87.3476
6.1145 | -87.3495
6.1148 | -87.3519
6.1151 | -87.3539
6.1155 | -87.3554
6.1158 | -87.3568
6.1161 | -87.3581
6.1164 | -87.3593
6.1167 | -87.3603
6.117 | -87.3613
6.1173 | -87.3622
6.1176 | -87.3633
6.1179 | -87.3651
6.1182 | -87.3672
6.1185 | -87.3686
6.1189 | -87.3696
6.1192 | -87.3705
6.1195 | -87.3716
6.1198 | -87.3721
6.1201 | -87.3731
6.1204 | -87.3737
6.1207 | -87.3749
6.121 | -87.3755
6.1213 | -87.3768
6.1216 | -87.3782
6.1219 | -87.379
6.1222 | -87.3801
6.1226 | -87.3807
6.1229 | -87.3815
6.1232 | -87.3825
6.1235 | -87.3838
6.1238 | -87.3861
6.1241 | -87.3877
6.1244 | -87.3886
6.1247 | -87.3896
6.125 | -87.3907
6.1253 | -87.3924
6.1256 | -87.3934
6.126 | -87.394
6.1263 | -87.395
6.1266 | -87.3958
6.1269 | -87.3967
6.1272 | -87.3979
6.1275 | -87.3987
6.1278 | -87.3995
6.1281 | -87.4007
6.1284 | -87.4015
6.1287 | -87.4025
6.129 | -87.4014
6.1293 | -87.4029
6.1297 | -87.4042
6.13 | -87.4049
6.1303 | -87.4054
6.1306 | -87.4057
6.1309 | -87.4068
6.1312 | -87.4084
6.1315 | -87.4095
6.1318 | -87.4099
6.1321 | -87.4107
6.1324 | -87.4118
6.1327 | -87.4128
6.1331 | -87.4137
6.1334 | -87.4146
6.1337 | -87.4152
6.134 | -87.4158
6.1343 | -87.4164
6.1346 | -87.4169
6.1349 | -87.418
6.1352 | -87.4187
6.1355 | -87.4187
6.1358 | -87.4194
6.1361 | -87.4197
6.1365 | -87.4193
6.1368 | -87.4201
6.1371 | -87.4212
6.1374 | -87.4222
6.1377 | -87.4227
6.138 | -87.4234
6.1383 | -87.4242
6.1386 | -87.4253
6.1389 | -87.4264
6.1392 | -87.4272
6.1395 | -87.4275
6.1398 | -87.4281
6.1402 | -87.4287
6.1405 | -87.4295
6.1408 | -87.4296
6.1411 | -87.4293
6.1414 | -87.43
6.1417 | -87.4315
6.142 | -87.4322
6.1423 | -87.4321
6.1426 | -87.432
6.1429 | -87.4318
6.1432 | -87.4327
6.1436 | -87.4346
6.1439 | -87.4362
6.1442 | -87.4373
6.1445 | -87.4379
6.1448 | -87.4383
6.1451 | -87.4394
6.1454 | -87.4407
6.1457 | -87.4413
6.146 | -87.4421
6.1463 | -87.4429
6.1466 | -87.4436
6.1469 | -87.444
6.1473 | -87.4444
6.1476 | -87.4453
6.1479 | -87.446
6.1482 | -87.4466
6.1485 | -87.4479
6.1488 | -87.4481
6.1491 | -87.4477
6.1494 | -87.4478
6.1497 | -87.4486
6.15 | -87.4494
6.1503 | -87.4505
6.1507 | -87.4532
6.151 | -87.4543
6.1513 | -87.4547
6.1516 | -87.4553
6.1519 | -87.457
6.1522 | -87.4588
6.1525 | -87.4606
6.1528 | -87.4618
6.1531 | -87.4619
6.1534 | -87.4613
6.1537 | -87.4616
6.154 | -87.4629
6.1544 | -87.4634
6.1547 | -87.4632
6.155 | -87.4637
6.1553 | -87.4649
6.1556 | -87.4655
6.1559 | -87.4657
6.1562 | -87.4656
6.1565 | -87.466
6.1568 | -87.4668
6.1571 | -87.4672
6.1574 | -87.4677
6.1578 | -87.4677
6.1581 | -87.4674
6.1584 | -87.4673
6.1587 | -87.4676
6.159 | -87.4683
6.1593 | -87.4686
6.1596 | -87.4681
6.1599 | -87.4681
6.1602 | -87.4684
6.1605 | -87.4688
6.1608 | -87.4694
6.1612 | -87.4693
6.1615 | -87.468
6.1618 | -87.4703
6.1621 | -87.4717
6.1624 | -87.4719
6.1627 | -87.4727
6.163 | -87.4734
6.1633 | -87.4738
6.1636 | -87.4746
6.1639 | -87.4747
6.1642 | -87.4754
6.1645 | -87.476
6.1649 | -87.476
6.1652 | -87.4753
6.1655 | -87.4754
6.1658 | -87.4762
6.1661 | -87.4768
6.1664 | -87.4778
6.1667 | -87.4794
6.167 | -87.4801
6.1673 | -87.4802
6.1676 | -87.4803
6.1679 | -87.4804
6.1683 | -87.4806
6.1686 | -87.4817
6.1689 | -87.4837
6.1692 | -87.4843
6.1695 | -87.484
6.1698 | -87.4833
6.1701 | -87.4833
6.1704 | -87.4841
6.1707 | -87.4852
6.171 | -87.4857
6.1713 | -87.4848
6.1716 | -87.485
6.172 | -87.4852
6.1723 | -87.4854
6.1726 | -87.4855
6.1729 | -87.4859
6.1732 | -87.4862
6.1735 | -87.4867
6.1738 | -87.4868
6.1741 | -87.4876
6.1744 | -87.4879
6.1747 | -87.4878
6.175 | -87.487
6.1754 | -87.487
6.1757 | -87.4874
6.176 | -87.4881
6.1763 | -87.4885
6.1766 | -87.4879
6.1769 | -87.4874
6.1772 | -87.4879
6.1775 | -87.4883
6.1778 | -87.4888
6.1781 | -87.4904
6.1784 | -87.4918
6.1788 | -87.4922
6.1791 | -87.4918
6.1794 | -87.4919
6.1797 | -87.4933
6.18 | -87.4946
6.1803 | -87.4958
6.1806 | -87.4961
6.1809 | -87.4961
6.1812 | -87.4963
6.1815 | -87.4969
6.1818 | -87.4991
6.1821 | -87.5
6.1825 | -87.4998
6.1828 | -87.5012
6.1831 | -87.502
6.1834 | -87.5024
6.1837 | -87.5021
6.184 | -87.5013
6.1843 | -87.5011
6.1846 | -87.5021
6.1849 | -87.5031
6.1852 | -87.5032
6.1855 | -87.5027
6.1859 | -87.503
6.1862 | -87.5034
6.1865 | -87.5045
6.1868 | -87.5056
6.1871 | -87.5059
6.1874 | -87.5048
6.1877 | -87.5043
6.188 | -87.5041
6.1883 | -87.5038
6.1886 | -87.5035
6.1889 | -87.5033
6.1893 | -87.5038
6.1896 | -87.5047
6.1899 | -87.5043
6.1902 | -87.5037
6.1905 | -87.5037
6.1908 | -87.5039
6.1911 | -87.5046
6.1914 | -87.505
6.1917 | -87.5066
6.192 | -87.5081
6.1923 | -87.509
6.1926 | -87.5093
6.193 | -87.5094
6.1933 | -87.5094
6.1936 | -87.5094
6.1939 | -87.5095
6.1942 | -87.5096
6.1945 | -87.5099
6.1948 | -87.5097
6.1951 | -87.5115
6.1954 | -87.5129
6.1957 | -87.5128
6.196 | -87.5129
6.1964 | -87.5128
6.1967 | -87.5128
6.197 | -87.5121
6.1973 | -87.5118
6.1976 | -87.512
6.1979 | -87.5117
6.1982 | -87.5114
6.1985 | -87.5113
6.1988 | -87.511
6.1991 | -87.5108
6.1994 | -87.5109
6.1997 | -87.51
6.2001 | -87.51
6.2004 | -87.5097
6.2007 | -87.5103
6.201 | -87.5111
6.2013 | -87.5118
6.2016 | -87.5119
6.2019 | -87.5119
6.2022 | -87.5126
6.2025 | -87.5126
6.2028 | -87.5125
6.2031 | -87.5125
6.2035 | -87.5124
6.2038 | -87.5117
6.2041 | -87.5115
6.2044 | -87.5113
6.2047 | -87.511
6.205 | -87.5122
6.2053 | -87.5136
6.2056 | -87.5138
6.2059 | -87.5133
6.2062 | -87.5129
6.2065 | -87.5131
6.2069 | -87.513
6.2072 | -87.5131
6.2075 | -87.5135
6.2078 | -87.5132
6.2081 | -87.5131
6.2084 | -87.5134
6.2087 | -87.5129
6.209 | -87.5117
6.2093 | -87.5114
6.2096 | -87.5112
6.2099 | -87.511
6.2102 | -87.5109
6.2106 | -87.5107
6.2109 | -87.5106
6.2112 | -87.5107
6.2115 | -87.5105
6.2118 | -87.5104
6.2121 | -87.5097
6.2124 | -87.5089
6.2127 | -87.5087
6.213 | -87.5079
6.2133 | -87.5075
6.2136 | -87.5074
6.214 | -87.5074
6.2143 | -87.5072
6.2146 | -87.5065
6.2149 | -87.5057
6.2152 | -87.505
6.2155 | -87.5051
6.2158 | -87.5054
6.2161 | -87.5058
6.2164 | -87.506
6.2167 | -87.5061
6.217 | -87.5059
6.2173 | -87.5041
6.2177 | -87.5039
6.218 | -87.5036
6.2183 | -87.5036
6.2186 | -87.504
6.2189 | -87.5036
6.2192 | -87.5029
6.2195 | -87.5029
6.2198 | -87.5018
6.2201 | -87.5018
6.2204 | -87.5025
6.2207 | -87.503
6.2211 | -87.503
6.2214 | -87.5029
6.2217 | -87.5021
6.222 | -87.5002
6.2223 | -87.4998
6.2226 | -87.4992
6.2229 | -87.4982
6.2232 | -87.4973
6.2235 | -87.4976
6.2238 | -87.4977
6.2241 | -87.4975
6.2244 | -87.4978
6.2248 | -87.4977
6.2251 | -87.4976
6.2254 | -87.4973
6.2257 | -87.4972
6.226 | -87.4966
6.2263 | -87.4958
6.2266 | -87.4954
6.2269 | -87.4944
6.2272 | -87.494
6.2275 | -87.493
6.2278 | -87.493
6.2282 | -87.4935
6.2285 | -87.4929
6.2288 | -87.4928
6.2291 | -87.4927
6.2294 | -87.4926
6.2297 | -87.4924
6.23 | -87.4922
6.2303 | -87.4922
6.2306 | -87.4923
6.2309 | -87.4921
6.2312 | -87.4914
6.2316 | -87.4903
6.2319 | -87.4892
6.2322 | -87.4887
6.2325 | -87.4891
6.2328 | -87.4889
6.2331 | -87.488
6.2334 | -87.4876
6.2337 | -87.4883
6.234 | -87.4895
6.2343 | -87.4891
6.2346 | -87.4889
6.2349 | -87.4887
6.2353 | -87.4879
6.2356 | -87.4874
6.2359 | -87.4874
6.2362 | -87.4877
6.2365 | -87.4873
6.2368 | -87.4868
6.2371 | -87.4865
6.2374 | -87.4869
6.2377 | -87.4867
6.238 | -87.4873
6.2383 | -87.4879
6.2387 | -87.488
6.239 | -87.4878
6.2393 | -87.4869
6.2396 | -87.4866
6.2399 | -87.4862
6.2402 | -87.486
6.2405 | -87.4858
6.2408 | -87.4856
6.2411 | -87.4851
6.2414 | -87.4836
6.2417 | -87.4827
6.242 | -87.4809
6.2424 | -87.4796
6.2427 | -87.4792
6.243 | -87.4801
6.2433 | -87.4805
6.2436 | -87.4788
6.2439 | -87.4773
6.2442 | -87.4762
6.2445 | -87.4753
6.2448 | -87.4738
6.2451 | -87.4737
6.2454 | -87.4733
6.2458 | -87.4721
6.2461 | -87.4708
6.2464 | -87.4699
6.2467 | -87.4692
6.247 | -87.4692
6.2473 | -87.4695
6.2476 | -87.4689
6.2479 | -87.469
6.2482 | -87.4687
6.2485 | -87.4681
6.2488 | -87.467
6.2492 | -87.4658
6.2495 | -87.4648
6.2498 | -87.4638
6.2501 | -87.4633
6.2504 | -87.4627
6.2507 | -87.4628
6.251 | -87.4621
6.2513 | -87.46
6.2516 | -87.4587
6.2519 | -87.4581
6.2522 | -87.4564
6.2525 | -87.4543
6.2529 | -87.4532
6.2532 | -87.4526
6.2535 | -87.4503
6.2538 | -87.4479
6.2541 | -87.4462
6.2544 | -87.4459
6.2547 | -87.4453
6.255 | -87.4446
6.2553 | -87.4438
6.2556 | -87.4445
6.2559 | -87.4454
6.2563 | -87.4445
6.2566 | -87.4431
6.2569 | -87.4425
6.2572 | -87.4417
6.2575 | -87.4406
6.2578 | -87.44
6.2581 | -87.4386
6.2584 | -87.4376
6.2587 | -87.4371
6.259 | -87.4364
6.2593 | -87.4356
6.2596 | -87.435
6.26 | -87.4344
6.2603 | -87.4337
6.2606 | -87.4326
6.2609 | -87.4319
6.2612 | -87.4317
6.2615 | -87.432
6.2618 | -87.4311
6.2621 | -87.4304
6.2624 | -87.4292
6.2627 | -87.4285
6.263 | -87.4283
6.2634 | -87.4267
6.2637 | -87.4244
6.264 | -87.4236
6.2643 | -87.4237
6.2646 | -87.4229
6.2649 | -87.4227
6.2652 | -87.422
6.2655 | -87.4199
6.2658 | -87.4173
6.2661 | -87.4156
6.2664 | -87.4139
6.2668 | -87.4133
6.2671 | -87.4127
6.2674 | -87.4119
6.2677 | -87.4114
6.268 | -87.4103
6.2683 | -87.4092
6.2686 | -87.4085
6.2689 | -87.4076
6.2692 | -87.4062
6.2695 | -87.4055
6.2698 | -87.4052
6.2701 | -87.4046
6.2705 | -87.4031
6.2708 | -87.4018
6.2711 | -87.4009
6.2714 | -87.3999
6.2717 | -87.3994
6.272 | -87.3986
6.2723 | -87.3972
6.2726 | -87.3962
6.2729 | -87.3954
6.2732 | -87.3946
6.2735 | -87.3952
6.2739 | -87.3943
6.2742 | -87.3937
6.2745 | -87.3929
6.2748 | -87.3925
6.2751 | -87.3917
6.2754 | -87.3916
6.2757 | -87.3913
6.276 | -87.3903
6.2763 | -87.3896
6.2766 | -87.3893
6.2769 | -87.3875
6.2772 | -87.3843
6.2776 | -87.383
6.2779 | -87.3828
6.2782 | -87.3823
6.2785 | -87.3817
6.2788 | -87.3812
6.2791 | -87.3805
6.2794 | -87.3793
6.2797 | -87.3778
6.28 | -87.3773
6.2803 | -87.3764
6.2806 | -87.3755
6.281 | -87.3734
6.2813 | -87.3716
6.2816 | -87.3707
6.2819 | -87.3696
6.2822 | -87.3686
6.2825 | -87.368
6.2828 | -87.3674
6.2831 | -87.3666
6.2834 | -87.3652
6.2837 | -87.3638
6.284 | -87.363
6.2843 | -87.3617
6.2847 | -87.3608
6.285 | -87.3597
6.2853 | -87.3585
6.2856 | -87.3578
6.2859 | -87.3573
6.2862 | -87.3564
6.2865 | -87.3553
6.2868 | -87.3549
6.2871 | -87.3543
6.2874 | -87.3534
6.2877 | -87.3529
6.2881 | -87.3525
6.2884 | -87.3514
6.2887 | -87.35
6.289 | -87.3475
6.2893 | -87.3457
6.2896 | -87.345
6.2899 | -87.3435
6.2902 | -87.3428
6.2905 | -87.3409
6.2908 | -87.3372
6.2911 | -87.3317
6.2915 | -87.3298
6.2918 | -87.3292
6.2921 | -87.3279
6.2924 | -87.3264
6.2927 | -87.3246
6.293 | -87.3223
6.2933 | -87.3201
6.2936 | -87.3183
6.2939 | -87.3164
6.2942 | -87.315
6.2945 | -87.3137
6.2948 | -87.3127
6.2952 | -87.3122
6.2955 | -87.3114
6.2958 | -87.3109
6.2961 | -87.3092
6.2964 | -87.3075
6.2967 | -87.3069
6.297 | -87.3061
6.2973 | -87.3039
6.2976 | -87.3017
6.2979 | -87.3005
6.2982 | -87.2987
6.2986 | -87.298
6.2989 | -87.297
6.2992 | -87.2955
6.2995 | -87.2941
6.2998 | -87.2925
6.3001 | -87.2891
6.3004 | -87.2862
6.3007 | -87.2845
6.301 | -87.2834
6.3013 | -87.2822
6.3016 | -87.2809
6.3019 | -87.2796
6.3023 | -87.2783
6.3026 | -87.2771
6.3029 | -87.2759
6.3032 | -87.2746
6.3035 | -87.2731
6.3038 | -87.2717
6.3041 | -87.2703
6.3044 | -87.2694
6.3047 | -87.2685
6.305 | -87.2676
6.3053 | -87.2666
6.3057 | -87.2653
6.306 | -87.2644
6.3063 | -87.2623
6.3066 | -87.2603
6.3069 | -87.2591
6.3072 | -87.2579
6.3075 | -87.2566
6.3078 | -87.2553
6.3081 | -87.2543
6.3084 | -87.2532
6.3087 | -87.2521
6.3091 | -87.2513
6.3094 | -87.2498
6.3097 | -87.2485
6.31 | -87.2465
6.3103 | -87.2433
6.3106 | -87.24
6.3109 | -87.2368
6.3112 | -87.234
6.3115 | -87.2322
6.3118 | -87.2311
6.3121 | -87.2303
6.3124 | -87.229
6.3128 | -87.2275
6.3131 | -87.226
6.3134 | -87.2246
6.3137 | -87.2226
6.314 | -87.2201
6.3143 | -87.2187
6.3146 | -87.2178
6.3149 | -87.2156
6.3152 | -87.2135
6.3155 | -87.2121
6.3158 | -87.2101
6.3162 | -87.2083
6.3165 | -87.2073
6.3168 | -87.2062
6.3171 | -87.2049
6.3174 | -87.2024
6.3177 | -87.1998
6.318 | -87.1973
6.3183 | -87.1951
6.3186 | -87.1936
6.3189 | -87.192
6.3192 | -87.1903
6.3195 | -87.1889
6.3199 | -87.1873
6.3202 | -87.1859
6.3205 | -87.1847
6.3208 | -87.183
6.3211 | -87.1823
6.3214 | -87.1824
6.3217 | -87.1812
6.322 | -87.1798
6.3223 | -87.1787
6.3226 | -87.177
6.3229 | -87.1744
6.3233 | -87.1729
6.3236 | -87.1721
6.3239 | -87.1706
6.3242 | -87.1683
6.3245 | -87.1654
6.3248 | -87.1636
6.3251 | -87.162
6.3254 | -87.1604
6.3257 | -87.1591
6.326 | -87.1573
6.3263 | -87.1548
6.3267 | -87.1524
6.327 | -87.1508
6.3273 | -87.1491
6.3276 | -87.1477
6.3279 | -87.1464
6.3282 | -87.1444
6.3285 | -87.142
6.3288 | -87.1399
6.3291 | -87.1376
6.3294 | -87.1351
6.3297 | -87.133
6.33 | -87.1314
6.3304 | -87.1298
6.3307 | -87.1283
6.331 | -87.1259
6.3313 | -87.1245
6.3316 | -87.1222
6.3319 | -87.1217
6.3322 | -87.1209
6.3325 | -87.1198
6.3328 | -87.1189
6.3331 | -87.1167
6.3334 | -87.1144
6.3338 | -87.1126
6.3341 | -87.1099
6.3344 | -87.1083
6.3347 | -87.1065
6.335 | -87.1047
6.3353 | -87.1028
6.3356 | -87.1007
6.3359 | -87.099
6.3362 | -87.0977
6.3365 | -87.097
6.3368 | -87.0959
6.3372 | -87.0946
6.3375 | -87.0928
6.3378 | -87.0915
6.3381 | -87.0902
6.3384 | -87.0894
6.3387 | -87.0885
6.339 | -87.0862
6.3393 | -87.0843
6.3396 | -87.0826
6.3399 | -87.0807
6.3402 | -87.0784
6.3405 | -87.0769
6.3409 | -87.0755
6.3412 | -87.0724
6.3415 | -87.0698
6.3418 | -87.0673
6.3421 | -87.0655
6.3424 | -87.0637
6.3427 | -87.0613
6.343 | -87.0592
6.3433 | -87.0573
6.3436 | -87.0555
6.3439 | -87.0538
6.3443 | -87.0521
6.3446 | -87.0513
6.3449 | -87.0497
6.3452 | -87.048
6.3455 | -87.0462
6.3458 | -87.0445
6.3461 | -87.0439
6.3464 | -87.0428
6.3467 | -87.0407
6.347 | -87.0395
6.3473 | -87.0376
6.3476 | -87.0351
6.348 | -87.033
6.3483 | -87.0312
6.3486 | -87.0295
6.3489 | -87.028
6.3492 | -87.0261
6.3495 | -87.0241
6.3498 | -87.0225
6.3501 | -87.0213
6.3504 | -87.0201
6.3507 | -87.0188
6.351 | -87.0171
6.3514 | -87.0161
6.3517 | -87.0122
6.352 | -87.0087
6.3523 | -87.0061
6.3526 | -87.0047
6.3529 | -87.003
6.3532 | -87.0013
6.3535 | -86.9999
6.3538 | -86.9976
6.3541 | -86.995
6.3544 | -86.9925
6.3547 | -86.9904
6.3551 | -86.9888
6.3554 | -86.9868
6.3557 | -86.985
6.356 | -86.9831
6.3563 | -86.9811
6.3566 | -86.9792
6.3569 | -86.9773
6.3572 | -86.9747
6.3575 | -86.9719
6.3578 | -86.9694
6.3581 | -86.9677
6.3585 | -86.9654
6.3588 | -86.9632
6.3591 | -86.9614
6.3594 | -86.9592
6.3597 | -86.9561
6.36 | -86.9531
6.3603 | -86.9513
6.3606 | -86.9474
6.3609 | -86.9453
6.3612 | -86.9434
6.3615 | -86.9409
6.3619 | -86.9382
6.3622 | -86.9356
6.3625 | -86.9332
6.3628 | -86.9301
6.3631 | -86.9281
6.3634 | -86.9261
6.3637 | -86.9241
6.364 | -86.9218
6.3643 | -86.92
6.3646 | -86.9178
6.3649 | -86.9151
6.3652 | -86.9133
6.3656 | -86.9117
6.3659 | -86.9096
6.3662 | -86.9062
6.3665 | -86.903
6.3668 | -86.9009
6.3671 | -86.8985
6.3674 | -86.8967
6.3677 | -86.8947
6.368 | -86.8925
6.3683 | -86.8901
6.3686 | -86.888
6.369 | -86.886
6.3693 | -86.8841
6.3696 | -86.8816
6.3699 | -86.8796
6.3702 | -86.8776
6.3705 | -86.8753
6.3708 | -86.873
6.3711 | -86.8704
6.3714 | -86.8684
6.3717 | -86.8666
6.372 | -86.8646
6.3723 | -86.8627
6.3727 | -86.8592
6.373 | -86.856
6.3733 | -86.8539
6.3736 | -86.8521
6.3739 | -86.8499
6.3742 | -86.8479
6.3745 | -86.8457
6.3748 | -86.8435
6.3751 | -86.8415
6.3754 | -86.8395
6.3757 | -86.8376
6.3761 | -86.8352
6.3764 | -86.8328
6.3767 | -86.8309
6.377 | -86.8293
6.3773 | -86.8273
6.3776 | -86.8247
6.3779 | -86.8226
6.3782 | -86.8206
6.3785 | -86.8184
6.3788 | -86.8164
6.3791 | -86.8145
6.3795 | -86.8128
6.3798 | -86.8111
6.3801 | -86.8093
6.3804 | -86.8078
6.3807 | -86.8065
6.381 | -86.8051
6.3813 | -86.8038
6.3816 | -86.8016
6.3819 | -86.7995
6.3822 | -86.7976
6.3825 | -86.7956
6.3828 | -86.7935
6.3832 | -86.7916
6.3835 | -86.7898
6.3838 | -86.7876
6.3841 | -86.7855
6.3844 | -86.7834
6.3847 | -86.7816
6.385 | -86.7803
6.3853 | -86.7778
6.3856 | -86.7759
6.3859 | -86.7736
6.3862 | -86.771
6.3866 | -86.7687
6.3869 | -86.7664
6.3872 | -86.7642
6.3875 | -86.7622
6.3878 | -86.7598
6.3881 | -86.7582
6.3884 | -86.7567
6.3887 | -86.7551
6.389 | -86.753
6.3893 | -86.7505
6.3896 | -86.748
6.3899 | -86.7458
6.3903 | -86.7447
6.3906 | -86.7434
6.3909 | -86.7416
6.3912 | -86.7401
6.3915 | -86.7377
6.3918 | -86.7345
6.3921 | -86.7322
6.3924 | -86.7298
6.3927 | -86.7274
6.393 | -86.725
6.3933 | -86.7229
6.3937 | -86.7207
6.394 | -86.7186
6.3943 | -86.7165
6.3946 | -86.7145
6.3949 | -86.7126
6.3952 | -86.7103
6.3955 | -86.7073
6.3958 | -86.7053
6.3961 | -86.7028
6.3964 | -86.7009
6.3967 | -86.6991
6.3971 | -86.6972
6.3974 | -86.6957
6.3977 | -86.6939
6.398 | -86.6919
6.3983 | -86.6899
6.3986 | -86.688
6.3989 | -86.6858
6.3992 | -86.684
6.3995 | -86.6814
6.3998 | -86.6797
6.4001 | -86.6779
6.4004 | -86.6751
6.4008 | -86.6728
6.4011 | -86.6705
6.4014 | -86.6683
6.4017 | -86.6653
6.402 | -86.6637
6.4023 | -86.6625
6.4026 | -86.6606
6.4029 | -86.6579
6.4032 | -86.6561
6.4035 | -86.6548
6.4038 | -86.6529
6.4042 | -86.6513
6.4045 | -86.6499
6.4048 | -86.6482
6.4051 | -86.6468
6.4054 | -86.6447
6.4057 | -86.6424
6.406 | -86.6405
6.4063 | -86.6381
6.4066 | -86.6363
6.4069 | -86.6345
6.4072 | -86.6326
6.4075 | -86.6305
6.4079 | -86.6288
6.4082 | -86.6279
6.4085 | -86.6267
6.4088 | -86.6249
6.4091 | -86.6227
6.4094 | -86.6208
6.4097 | -86.6192
6.41 | -86.6173
6.4103 | -86.6154
6.4106 | -86.6137
6.4109 | -86.6125
6.4113 | -86.611
6.4116 | -86.6088
6.4119 | -86.6066
6.4122 | -86.6044
6.4125 | -86.6022
6.4128 | -86.5996
6.4131 | -86.596
6.4134 | -86.593
6.4137 | -86.5912
6.414 | -86.5893
6.4143 | -86.5869
6.4146 | -86.5848
6.415 | -86.5829
6.4153 | -86.5811
6.4156 | -86.5792
6.4159 | -86.577
6.4162 | -86.5746
6.4165 | -86.572
6.4168 | -86.5699
6.4171 | -86.5677
6.4174 | -86.5659
6.4177 | -86.565
6.418 | -86.5635
6.4184 | -86.5614
6.4187 | -86.5592
6.419 | -86.5571
6.4193 | -86.5555
6.4196 | -86.5543
6.4199 | -86.5528
6.4202 | -86.5509
6.4205 | -86.5492
6.4208 | -86.5472
6.4211 | -86.5457
6.4214 | -86.544
6.4218 | -86.5419
6.4221 | -86.5396
6.4224 | -86.5375
6.4227 | -86.5363
6.423 | -86.5333
6.4233 | -86.5315
6.4236 | -86.5303
6.4239 | -86.5289
6.4242 | -86.5269
6.4245 | -86.5239
6.4248 | -86.5215
6.4251 | -86.5193
6.4255 | -86.5167
6.4258 | -86.5142
6.4261 | -86.5123
6.4264 | -86.5093
6.4267 | -86.5062
6.427 | -86.5043
6.4273 | -86.5025
6.4276 | -86.5004
6.4279 | -86.4981
6.4282 | -86.4961
6.4285 | -86.4944
6.4289 | -86.4921
6.4292 | -86.4898
6.4295 | -86.4878
6.4298 | -86.4854
6.4301 | -86.4829
6.4304 | -86.4801
6.4307 | -86.4774
6.431 | -86.475
6.4313 | -86.4729
6.4316 | -86.4714
6.4319 | -86.4701
6.4322 | -86.4686
6.4326 | -86.4666
6.4329 | -86.4647
6.4332 | -86.4636
6.4335 | -86.4625
6.4338 | -86.4609
6.4341 | -86.4591
6.4344 | -86.4575
6.4347 | -86.4554
6.435 | -86.4533
6.4353 | -86.4513
6.4356 | -86.4496
6.436 | -86.4478
6.4363 | -86.4462
6.4366 | -86.4446
6.4369 | -86.4432
6.4372 | -86.4416
6.4375 | -86.4399
6.4378 | -86.4379
6.4381 | -86.4357
6.4384 | -86.4337
6.4387 | -86.4321
6.439 | -86.4303
6.4394 | -86.4281
6.4397 | -86.4262
6.44 | -86.4244
6.4403 | -86.4216
6.4406 | -86.4196
6.4409 | -86.4184
6.4412 | -86.4166
6.4415 | -86.4144
6.4418 | -86.4125
6.4421 | -86.4104
6.4424 | -86.408
6.4427 | -86.4056
6.4431 | -86.4033
6.4434 | -86.4016
6.4437 | -86.3996
6.444 | -86.3978
6.4443 | -86.3956
6.4446 | -86.3931
6.4449 | -86.3911
6.4452 | -86.3893
6.4455 | -86.3874
6.4458 | -86.3859
6.4461 | -86.3839
6.4465 | -86.3802
6.4468 | -86.3774
6.4471 | -86.3756
6.4474 | -86.3738
6.4477 | -86.3716
6.448 | -86.3696
6.4483 | -86.3676
6.4486 | -86.3663
6.4489 | -86.3655
6.4492 | -86.3644
6.4495 | -86.3632
6.4498 | -86.3607
6.4502 | -86.3589
6.4505 | -86.3567
6.4508 | -86.3546
6.4511 | -86.3529
6.4514 | -86.3514
6.4517 | -86.3497
6.452 | -86.348
6.4523 | -86.3458
6.4526 | -86.3433
6.4529 | -86.341
6.4532 | -86.3392
6.4536 | -86.3374
6.4539 | -86.3357
6.4542 | -86.3345
6.4545 | -86.3328
6.4548 | -86.3315
6.4551 | -86.3298
6.4554 | -86.3281
6.4557 | -86.3263
6.456 | -86.325
6.4563 | -86.3231
6.4566 | -86.3214
6.457 | -86.3195
6.4573 | -86.3167
6.4576 | -86.3149
6.4579 | -86.3126
6.4582 | -86.3103
6.4585 | -86.3086
6.4588 | -86.3076
6.4591 | -86.3064
6.4594 | -86.3047
6.4597 | -86.3034
6.46 | -86.3022
6.4603 | -86.3006
6.4607 | -86.2992
6.461 | -86.2983
6.4613 | -86.297
6.4616 | -86.2956
6.4619 | -86.2939
6.4622 | -86.2921
6.4625 | -86.2903
6.4628 | -86.2885
6.4631 | -86.2867
6.4634 | -86.2848
6.4637 | -86.2832
6.4641 | -86.2814
6.4644 | -86.2801
6.4647 | -86.2787
6.465 | -86.2772
6.4653 | -86.2752
6.4656 | -86.273
6.4659 | -86.2711
6.4662 | -86.2698
6.4665 | -86.2682
6.4668 | -86.2665
6.4671 | -86.2656
6.4675 | -86.2651
6.4678 | -86.2642
6.4681 | -86.2631
6.4684 | -86.2615
6.4687 | -86.2604
6.469 | -86.2597
6.4693 | -86.2582
6.4696 | -86.2569
6.4699 | -86.2559
6.4702 | -86.2545
6.4705 | -86.2532
6.4708 | -86.252
6.4712 | -86.2507
6.4715 | -86.2494
6.4718 | -86.2483
6.4721 | -86.2471
6.4724 | -86.2459
6.4727 | -86.2444
6.473 | -86.2425
6.4733 | -86.2411
6.4736 | -86.2398
6.4739 | -86.2384
6.4742 | -86.237
6.4746 | -86.236
6.4749 | -86.2351
6.4752 | -86.234
6.4755 | -86.2323
6.4758 | -86.2306
6.4761 | -86.2292
6.4764 | -86.2285
6.4767 | -86.2272
6.477 | -86.226
6.4773 | -86.2259
6.4776 | -86.2262
6.4779 | -86.2249
6.4783 | -86.2232
6.4786 | -86.2219
6.4789 | -86.2211
6.4792 | -86.2204
6.4795 | -86.2201
6.4798 | -86.2187
6.4801 | -86.2175
6.4804 | -86.2164
6.4807 | -86.2151
6.481 | -86.2137
6.4813 | -86.2125
6.4817 | -86.2117
6.482 | -86.2109
6.4823 | -86.2097
6.4826 | -86.2083
6.4829 | -86.2067
6.4832 | -86.2049
6.4835 | -86.2035
6.4838 | -86.2021
6.4841 | -86.2009
6.4844 | -86.2001
6.4847 | -86.1992
6.485 | -86.1984
6.4854 | -86.1979
6.4857 | -86.1967
6.486 | -86.1957
6.4863 | -86.1948
6.4866 | -86.1936
6.4869 | -86.1925
6.4872 | -86.1916
6.4875 | -86.1906
6.4878 | -86.1892
6.4881 | -86.1878
6.4884 | -86.1857
6.4888 | -86.1844
6.4891 | -86.1833
6.4894 | -86.1818
6.4897 | -86.1806
6.49 | -86.1797
6.4903 | -86.1786
6.4906 | -86.1774
6.4909 | -86.1763
6.4912 | -86.1744
6.4915 | -86.1725
6.4918 | -86.1713
6.4922 | -86.1703
6.4925 | -86.1693
6.4928 | -86.1684
6.4931 | -86.1673
6.4934 | -86.1663
6.4937 | -86.1647
6.494 | -86.1628
6.4943 | -86.1624
6.4946 | -86.1615
6.4949 | -86.1603
6.4952 | -86.1593
6.4955 | -86.1581
6.4959 | -86.1563
6.4962 | -86.1547
6.4965 | -86.1538
6.4968 | -86.1529
6.4971 | -86.1521
6.4974 | -86.1513
6.4977 | -86.1509
6.498 | -86.15
6.4983 | -86.1492
6.4986 | -86.1483
6.4989 | -86.1472
6.4993 | -86.1462
6.4996 | -86.1451
6.4999 | -86.1443
6.5002 | -86.1433
6.5005 | -86.1425
6.5008 | -86.1415
6.5011 | -86.1402
6.5014 | -86.1388
6.5017 | -86.1377
6.502 | -86.1362
6.5023 | -86.1349
6.5026 | -86.1343
6.503 | -86.1334
6.5033 | -86.1335
6.5036 | -86.1332
6.5039 | -86.1323
6.5042 | -86.1314
6.5045 | -86.1313
6.5048 | -86.1314
6.5051 | -86.1309
6.5054 | -86.13
6.5057 | -86.1293
6.506 | -86.129
6.5064 | -86.129
6.5067 | -86.1278
6.507 | -86.1268
6.5073 | -86.1259
6.5076 | -86.1248
6.5079 | -86.1228
6.5082 | -86.1203
6.5085 | -86.1182
6.5088 | -86.1173
6.5091 | -86.1167
6.5094 | -86.1168
6.5098 | -86.1168
6.5101 | -86.1164
6.5104 | -86.1153
6.5107 | -86.1146
6.511 | -86.1138
6.5113 | -86.1132
6.5116 | -86.1127
6.5119 | -86.1122
6.5122 | -86.1123
6.5125 | -86.1122
6.5128 | -86.1121
6.5131 | -86.1122
6.5135 | -86.1116
6.5138 | -86.1118
6.5141 | -86.1129
6.5144 | -86.1128
6.5147 | -86.1121
6.515 | -86.1114
6.5153 | -86.1105
6.5156 | -86.1101
6.5159 | -86.1096
6.5162 | -86.1087
6.5165 | -86.1081
6.5169 | -86.1074
6.5172 | -86.1067
6.5175 | -86.1057
6.5178 | -86.1049
6.5181 | -86.103
6.5184 | -86.1017
6.5187 | -86.1011
6.519 | -86.1011
6.5193 | -86.1011
6.5196 | -86.1008
6.5199 | -86.1004
6.5202 | -86.0998
6.5206 | -86.0984
6.5209 | -86.0984
6.5212 | -86.098
6.5215 | -86.098
6.5218 | -86.0978
6.5221 | -86.0973
6.5224 | -86.0969
6.5227 | -86.0966
6.523 | -86.096
6.5233 | -86.0953
6.5236 | -86.0947
6.524 | -86.0943
6.5243 | -86.0936
6.5246 | -86.0933
6.5249 | -86.0928
6.5252 | -86.0919
6.5255 | -86.0916
6.5258 | -86.0915
6.5261 | -86.0909
6.5264 | -86.0904
6.5267 | -86.0896
6.527 | -86.0885
6.5274 | -86.0877
6.5277 | -86.0876
6.528 | -86.088
6.5283 | -86.0873
6.5286 | -86.0869
6.5289 | -86.0863
6.5292 | -86.0851
6.5295 | -86.0842
6.5298 | -86.0849
6.5301 | -86.0844
6.5304 | -86.0845
6.5307 | -86.0842
6.5311 | -86.0841
6.5314 | -86.0839
6.5317 | -86.0836
6.532 | -86.0832
6.5323 | -86.0826
6.5326 | -86.0822
6.5329 | -86.0819
6.5332 | -86.0817
6.5335 | -86.0814
6.5338 | -86.0811
6.5341 | -86.0809
6.5345 | -86.0808
6.5348 | -86.0805
6.5351 | -86.0799
6.5354 | -86.079
6.5357 | -86.0774
6.536 | -86.0759
6.5363 | -86.0751
6.5366 | -86.0748
6.5369 | -86.075
6.5372 | -86.0756
6.5375 | -86.0756
6.5378 | -86.0752
6.5382 | -86.0737
6.5385 | -86.0731
6.5388 | -86.0738
6.5391 | -86.0736
6.5394 | -86.0735
6.5397 | -86.0734
6.54 | -86.0735
6.5403 | -86.0732
6.5406 | -86.0732
6.5409 | -86.0732
6.5412 | -86.073
6.5416 | -86.0728
6.5419 | -86.0729
6.5422 | -86.073
6.5425 | -86.0728
6.5428 | -86.0725
6.5431 | -86.0724
6.5434 | -86.0738
6.5437 | -86.0759
6.544 | -86.0764
6.5443 | -86.0763
6.5446 | -86.0762
6.545 | -86.076
6.5453 | -86.0756
6.5456 | -86.0747
6.5459 | -86.0743
6.5462 | -86.0741
6.5465 | -86.0735
6.5468 | -86.0729
6.5471 | -86.0733
6.5474 | -86.0737
6.5477 | -86.0745
6.548 | -86.0745
6.5483 | -86.0745
6.5487 | -86.0745
6.549 | -86.0745
6.5493 | -86.074
6.5496 | -86.0739
6.5499 | -86.0735
6.5502 | -86.0738
6.5505 | -86.0737
6.5508 | -86.0733
6.5511 | -86.0733
6.5514 | -86.0722
6.5517 | -86.071
6.5521 | -86.0703
6.5524 | -86.0698
6.5527 | -86.0694
6.553 | -86.0696
6.5533 | -86.0698
6.5536 | -86.0701
6.5539 | -86.0702
6.5542 | -86.0703
6.5545 | -86.0703
6.5548 | -86.0704
6.5551 | -86.0705
6.5554 | -86.0705
6.5558 | -86.0705
6.5561 | -86.0701
6.5564 | -86.0695
6.5567 | -86.0707
6.557 | -86.0706
6.5573 | -86.0707
6.5576 | -86.0704
6.5579 | -86.0703
6.5582 | -86.0701
6.5585 | -86.0702
6.5588 | -86.0704
6.5592 | -86.0703
6.5595 | -86.07
6.5598 | -86.0704
6.5601 | -86.0705
6.5604 | -86.0707
6.5607 | -86.0709
6.561 | -86.0711
6.5613 | -86.0716
6.5616 | -86.0711
6.5619 | -86.0706
6.5622 | -86.0708
6.5625 | -86.0704
6.5629 | -86.0705
6.5632 | -86.0708
6.5635 | -86.0708
6.5638 | -86.0705
6.5641 | -86.0702
6.5644 | -86.07
6.5647 | -86.0695
6.565 | -86.069
6.5653 | -86.0695
6.5656 | -86.0701
6.5659 | -86.0701
6.5663 | -86.07
6.5666 | -86.0705
6.5669 | -86.0706
6.5672 | -86.0707
6.5675 | -86.071
6.5678 | -86.0711
6.5681 | -86.0711
6.5684 | -86.0715
6.5687 | -86.0717
6.569 | -86.0715
6.5693 | -86.0712
6.5697 | -86.0721
6.57 | -86.0734
6.5703 | -86.0736
6.5706 | -86.0738
6.5709 | -86.0738
6.5712 | -86.0742
6.5715 | -86.0748
6.5718 | -86.0752
6.5721 | -86.0756
6.5724 | -86.0762
6.5727 | -86.0763
6.573 | -86.0762
6.5734 | -86.0757
6.5737 | -86.0754
6.574 | -86.0752
6.5743 | -86.0753
6.5746 | -86.076
6.5749 | -86.0765
6.5752 | -86.077
6.5755 | -86.0771
6.5758 | -86.0777
6.5761 | -86.0781
6.5764 | -86.0788
6.5768 | -86.0798
6.5771 | -86.0796
6.5774 | -86.0802
6.5777 | -86.0814
6.578 | -86.0815
6.5783 | -86.0815
6.5786 | -86.0819
6.5789 | -86.0822
6.5792 | -86.0826
6.5795 | -86.083
6.5798 | -86.0832
6.5801 | -86.0829
6.5805 | -86.0829
6.5808 | -86.0833
6.5811 | -86.0834
6.5814 | -86.0833
6.5817 | -86.0838
6.582 | -86.0841
6.5823 | -86.0848
6.5826 | -86.0852
6.5829 | -86.0862
6.5832 | -86.0864
6.5835 | -86.0873
6.5839 | -86.0882
6.5842 | -86.0886
6.5845 | -86.0894
6.5848 | -86.0899
6.5851 | -86.0901
6.5854 | -86.0902
6.5857 | -86.0906
6.586 | -86.0909
6.5863 | -86.0913
6.5866 | -86.0918
6.5869 | -86.0919
6.5873 | -86.0923
6.5876 | -86.0926
6.5879 | -86.0929
6.5882 | -86.0933
6.5885 | -86.0934
6.5888 | -86.0927
6.5891 | -86.093
6.5894 | -86.0935
6.5897 | -86.0941
6.59 | -86.0945
6.5903 | -86.0951
6.5906 | -86.095
6.591 | -86.0951
6.5913 | -86.0965
6.5916 | -86.0974
6.5919 | -86.0978
6.5922 | -86.099
6.5925 | -86.0995
6.5928 | -86.1003
6.5931 | -86.1008
6.5934 | -86.1014
6.5937 | -86.1015
6.594 | -86.1005
6.5944 | -86.1008
6.5947 | -86.1014
6.595 | -86.1015
6.5953 | -86.1018
6.5956 | -86.1027
6.5959 | -86.1041
6.5962 | -86.1047
6.5965 | -86.1052
6.5968 | -86.1061
6.5971 | -86.1069
6.5974 | -86.1075
6.5977 | -86.1081
6.5981 | -86.1086
6.5984 | -86.1086
6.5987 | -86.1091
6.599 | -86.1101
6.5993 | -86.1108
6.5996 | -86.1113
6.5999 | -86.1119
6.6002 | -86.1126
6.6005 | -86.1138
6.6008 | -86.1153
6.6011 | -86.1165
6.6015 | -86.1176
6.6018 | -86.1182
6.6021 | -86.1187
6.6024 | -86.1193
6.6027 | -86.1205
6.603 | -86.1215
6.6033 | -86.1223
6.6036 | -86.1228
6.6039 | -86.1218
6.6042 | -86.1218
6.6045 | -86.1236
6.6049 | -86.1247
6.6052 | -86.1253
6.6055 | -86.1251
6.6058 | -86.1238
6.6061 | -86.1235
6.6064 | -86.1241
6.6067 | -86.125
6.607 | -86.1257
6.6073 | -86.1263
6.6076 | -86.1268
6.6079 | -86.1269
6.6082 | -86.1272
6.6086 | -86.1274
6.6089 | -86.1264
6.6092 | -86.126
6.6095 | -86.1263
6.6098 | -86.1266
6.6101 | -86.1276
6.6104 | -86.128
6.6107 | -86.1283
6.611 | -86.129
6.6113 | -86.1292
6.6116 | -86.129
6.612 | -86.129
6.6123 | -86.1291
6.6126 | -86.1296
6.6129 | -86.1306
6.6132 | -86.1315
6.6135 | -86.1327
6.6138 | -86.1333
6.6141 | -86.1336
6.6144 | -86.1338
6.6147 | -86.1349
6.615 | -86.1367
6.6153 | -86.1376
6.6157 | -86.1382
6.616 | -86.1385
6.6163 | -86.1389
6.6166 | -86.1391
6.6169 | -86.1395
6.6172 | -86.1399
6.6175 | -86.1401
6.6178 | -86.1408
6.6181 | -86.1411
6.6184 | -86.1414
6.6187 | -86.1422
6.6191 | -86.143
6.6194 | -86.144
6.6197 | -86.1446
6.62 | -86.145
6.6203 | -86.1454
6.6206 | -86.146
6.6209 | -86.1469
6.6212 | -86.1478
6.6215 | -86.1484
6.6218 | -86.1489
6.6221 | -86.1495
6.6225 | -86.1499
6.6228 | -86.15
6.6231 | -86.1502
6.6234 | -86.151
6.6237 | -86.1521
6.624 | -86.1527
6.6243 | -86.1533
6.6246 | -86.1538
6.6249 | -86.1546
6.6252 | -86.1554
6.6255 | -86.1561
6.6258 | -86.1563
6.6262 | -86.1561
6.6265 | -86.1566
6.6268 | -86.1572
6.6271 | -86.1577
6.6274 | -86.1572
6.6277 | -86.1564
6.628 | -86.1569
6.6283 | -86.1573
6.6286 | -86.158
6.6289 | -86.1588
6.6292 | -86.159
6.6296 | -86.1591
6.6299 | -86.1595
6.6302 | -86.1601
6.6305 | -86.1611
6.6308 | -86.1613
6.6311 | -86.1619
6.6314 | -86.1626
6.6317 | -86.1631
6.632 | -86.1635
6.6323 | -86.1646
6.6326 | -86.1658
6.6329 | -86.1665
6.6333 | -86.1674
6.6336 | -86.1684
6.6339 | -86.1689
6.6342 | -86.1692
6.6345 | -86.1701
6.6348 | -86.171
6.6351 | -86.1718
6.6354 | -86.172
6.6357 | -86.1705
6.636 | -86.1703
6.6363 | -86.1708
6.6367 | -86.1711
6.637 | -86.172
6.6373 | -86.1722
6.6376 | -86.1726
6.6379 | -86.1732
6.6382 | -86.1741
6.6385 | -86.1745
6.6388 | -86.1746
6.6391 | -86.1751
6.6394 | -86.1755
6.6397 | -86.1756
6.6401 | -86.1752
6.6404 | -86.175
6.6407 | -86.1753
6.641 | -86.1757
6.6413 | -86.1762
6.6416 | -86.1768
6.6419 | -86.1773
6.6422 | -86.1779
6.6425 | -86.1783
6.6428 | -86.1773
6.6431 | -86.1759
6.6434 | -86.1756
6.6438 | -86.1763
6.6441 | -86.1763
6.6444 | -86.1759
6.6447 | -86.1763
6.645 | -86.1769
6.6453 | -86.1779
6.6456 | -86.178
6.6459 | -86.1792
6.6462 | -86.1788
6.6465 | -86.1786
6.6468 | -86.179
6.6472 | -86.1794
6.6475 | -86.1799
6.6478 | -86.1804
6.6481 | -86.1809
6.6484 | -86.1809
6.6487 | -86.1797
6.649 | -86.1788
6.6493 | -86.1792
6.6496 | -86.1799
6.6499 | -86.1808
6.6502 | -86.1814
6.6505 | -86.1818
6.6509 | -86.1827
6.6512 | -86.1832
6.6515 | -86.1835
6.6518 | -86.1832
6.6521 | -86.1842
6.6524 | -86.1853
6.6527 | -86.186
6.653 | -86.1865
6.6533 | -86.1869
6.6536 | -86.1873
6.6539 | -86.1879
6.6543 | -86.1884
6.6546 | -86.1894
6.6549 | -86.1899
6.6552 | -86.1906
6.6555 | -86.1918
6.6558 | -86.1925
6.6561 | -86.1928
6.6564 | -86.1935
6.6567 | -86.1938
6.657 | -86.1943
6.6573 | -86.1948
6.6577 | -86.1955
6.658 | -86.1963
6.6583 | -86.1968
6.6586 | -86.197
6.6589 | -86.1975
6.6592 | -86.1986
6.6595 | -86.1994
6.6598 | -86.2003
6.6601 | -86.2007
6.6604 | -86.2006
6.6607 | -86.2001
6.661 | -86.2001
6.6614 | -86.2009
6.6617 | -86.202
6.662 | -86.2026
6.6623 | -86.2031
6.6626 | -86.2031
6.6629 | -86.2018
6.6632 | -86.2015
6.6635 | -86.2021
6.6638 | -86.2025
6.6641 | -86.2031
6.6644 | -86.2038
6.6648 | -86.2046
6.6651 | -86.2051
6.6654 | -86.2051
6.6657 | -86.2041
6.666 | -86.204
6.6663 | -86.2044
6.6666 | -86.2052
6.6669 | -86.2058
6.6672 | -86.207
6.6675 | -86.2075
6.6678 | -86.2076
6.6681 | -86.2079
6.6685 | -86.2084
6.6688 | -86.2091
6.6691 | -86.2097
6.6694 | -86.211
6.6697 | -86.2116
6.67 | -86.2123
6.6703 | -86.2132
6.6706 | -86.214
6.6709 | -86.2143
6.6712 | -86.2144
6.6715 | -86.2146
6.6719 | -86.2147
6.6722 | -86.2146
6.6725 | -86.2143
6.6728 | -86.2145
6.6731 | -86.2147
6.6734 | -86.2152
6.6737 | -86.2155
6.674 | -86.216
6.6743 | -86.2164
6.6746 | -86.2172
6.6749 | -86.2172
6.6753 | -86.2162
6.6756 | -86.2161
6.6759 | -86.2168
6.6762 | -86.2165
6.6765 | -86.2162
6.6768 | -86.2165
6.6771 | -86.2167
6.6774 | -86.2168
6.6777 | -86.217
6.678 | -86.2173
6.6783 | -86.2176
6.6786 | -86.2177
6.679 | -86.2179
6.6793 | -86.218
6.6796 | -86.2183
6.6799 | -86.2188
6.6802 | -86.2191
6.6805 | -86.2191
6.6808 | -86.2195
6.6811 | -86.2201
6.6814 | -86.2203
6.6817 | -86.2206
6.682 | -86.221
6.6824 | -86.2205
6.6827 | -86.22
6.683 | -86.2203
6.6833 | -86.2204
6.6836 | -86.2212
6.6839 | -86.222
6.6842 | -86.2221
6.6845 | -86.2221
6.6848 | -86.2225
6.6851 | -86.2227
6.6854 | -86.2231
6.6857 | -86.2235
6.6861 | -86.2234
6.6864 | -86.2235
6.6867 | -86.2237
6.687 | -86.2241
6.6873 | -86.2239
6.6876 | -86.2236
6.6879 | -86.2246
6.6882 | -86.225
6.6885 | -86.2254
6.6888 | -86.2256
6.6891 | -86.2253
6.6895 | -86.2254
6.6898 | -86.2253
6.6901 | -86.2253
6.6904 | -86.225
6.6907 | -86.2254
6.691 | -86.2265
6.6913 | -86.2274
6.6916 | -86.2277
6.6919 | -86.228
6.6922 | -86.2282
6.6925 | -86.2291
6.6928 | -86.23
6.6932 | -86.2307
6.6935 | -86.2312
6.6938 | -86.2319
6.6941 | -86.2327
6.6944 | -86.2329
6.6947 | -86.2331
6.695 | -86.2338
6.6953 | -86.2348
6.6956 | -86.2344
6.6959 | -86.2335
6.6962 | -86.2337
6.6966 | -86.2344
6.6969 | -86.235
6.6972 | -86.2351
6.6975 | -86.2352
6.6978 | -86.2356
6.6981 | -86.2358
6.6984 | -86.2359
6.6987 | -86.2368
6.699 | -86.2372
6.6993 | -86.2374
6.6996 | -86.2377
6.7 | -86.2374
6.7003 | -86.2371
6.7006 | -86.2374
6.7009 | -86.2385
6.7012 | -86.2388
6.7015 | -86.2387
6.7018 | -86.2392
6.7021 | -86.2404
6.7024 | -86.2411
6.7027 | -86.2413
6.703 | -86.2426
6.7033 | -86.2438
6.7037 | -86.2445
6.704 | -86.2451
6.7043 | -86.2446
6.7046 | -86.2445
6.7049 | -86.2447
6.7052 | -86.245
6.7055 | -86.2447
6.7058 | -86.2449
6.7061 | -86.2451
6.7064 | -86.2457
6.7067 | -86.2459
6.7071 | -86.2462
6.7074 | -86.2463
6.7077 | -86.2461
6.708 | -86.2464
6.7083 | -86.2471
6.7086 | -86.2475
6.7089 | -86.2477
6.7092 | -86.2483
6.7095 | -86.2486
6.7098 | -86.249
6.7101 | -86.2488
6.7104 | -86.2491
6.7108 | -86.2494
6.7111 | -86.2495
6.7114 | -86.2494
6.7117 | -86.2497
6.712 | -86.2498
6.7123 | -86.2502
6.7126 | -86.2506
6.7129 | -86.2501
6.7132 | -86.2505
6.7135 | -86.2507
6.7138 | -86.2518
6.7142 | -86.2532
6.7145 | -86.2546
6.7148 | -86.2548
6.7151 | -86.2545
6.7154 | -86.2541
6.7157 | -86.2535
6.716 | -86.2546
6.7163 | -86.2559
6.7166 | -86.2565
6.7169 | -86.2567
6.7172 | -86.2574
6.7176 | -86.2574
6.7179 | -86.2573
6.7182 | -86.2575
6.7185 | -86.2576
6.7188 | -86.2579
6.7191 | -86.2587
6.7194 | -86.2588
6.7197 | -86.2581
6.72 | -86.2577
6.7203 | -86.2575
6.7206 | -86.2574
6.7209 | -86.2573
6.7213 | -86.2578
6.7216 | -86.258
6.7219 | -86.2579
6.7222 | -86.258
6.7225 | -86.2579
6.7228 | -86.2578
6.7231 | -86.2573
6.7234 | -86.2569
6.7237 | -86.2567
6.724 | -86.2568
6.7243 | -86.257
6.7247 | -86.2569
6.725 | -86.2572
6.7253 | -86.2567
6.7256 | -86.2569
6.7259 | -86.2569
6.7262 | -86.2568
6.7265 | -86.2568
6.7268 | -86.2569
6.7271 | -86.2573
6.7274 | -86.2575
6.7277 | -86.2582
6.728 | -86.2582
6.7284 | -86.259
6.7287 | -86.2609
6.729 | -86.2615
6.7293 | -86.2614
6.7296 | -86.2608
6.7299 | -86.2607
6.7302 | -86.2609
6.7305 | -86.2608
6.7308 | -86.261
6.7311 | -86.2619
6.7314 | -86.2625
6.7318 | -86.2618
6.7321 | -86.261
6.7324 | -86.2612
6.7327 | -86.2621
6.733 | -86.2628
6.7333 | -86.2629
6.7336 | -86.263
6.7339 | -86.2627
6.7342 | -86.2623
6.7345 | -86.2624
6.7348 | -86.2623
6.7352 | -86.2623
6.7355 | -86.2627
6.7358 | -86.2633
6.7361 | -86.2632
6.7364 | -86.2633
6.7367 | -86.2633
6.737 | -86.2634
6.7373 | -86.2637
6.7376 | -86.2639
6.7379 | -86.2636
6.7382 | -86.2632
6.7385 | -86.2631
6.7389 | -86.2628
6.7392 | -86.2625
6.7395 | -86.2628
6.7398 | -86.263
6.7401 | -86.2628
6.7404 | -86.263
6.7407 | -86.2635
6.741 | -86.2642
6.7413 | -86.2644
6.7416 | -86.2642
6.7419 | -86.2644
6.7423 | -86.2647
6.7426 | -86.2646
6.7429 | -86.2654
6.7432 | -86.2673
6.7435 | -86.2705
6.7438 | -86.2708
6.7441 | -86.2719
6.7444 | -86.2717
6.7447 | -86.2718
6.745 | -86.272
6.7453 | -86.2726
6.7456 | -86.2735
6.746 | -86.2747
6.7463 | -86.2754
6.7466 | -86.2762
6.7469 | -86.2761
6.7472 | -86.2761
6.7475 | -86.2762
6.7478 | -86.2762
6.7481 | -86.2767
6.7484 | -86.2769
6.7487 | -86.277
6.749 | -86.2772
6.7494 | -86.2771
6.7497 | -86.2771
6.75 | -86.2778
6.7503 | -86.2786
6.7506 | -86.2789
6.7509 | -86.2791
6.7512 | -86.2795
6.7515 | -86.2796
6.7518 | -86.2793
6.7521 | -86.2794
6.7524 | -86.2795
6.7528 | -86.2798
6.7531 | -86.2794
6.7534 | -86.2804
6.7537 | -86.2816
6.754 | -86.282
6.7543 | -86.2821
6.7546 | -86.2822
6.7549 | -86.2823
6.7552 | -86.2822
6.7555 | -86.2824
6.7558 | -86.2824
6.7561 | -86.2824
6.7565 | -86.2825
6.7568 | -86.2821
6.7571 | -86.2816
6.7574 | -86.2815
6.7577 | -86.2815
6.758 | -86.2816
6.7583 | -86.2818
6.7586 | -86.2816
6.7589 | -86.2815
6.7592 | -86.2813
6.7595 | -86.282
6.7599 | -86.2825
6.7602 | -86.2825
6.7605 | -86.2824
6.7608 | -86.2822
6.7611 | -86.2808
6.7614 | -86.2796
6.7617 | -86.2801
6.762 | -86.2798
6.7623 | -86.2797
6.7626 | -86.2796
6.7629 | -86.2796
6.7632 | -86.2799
6.7636 | -86.2798
6.7639 | -86.2801
6.7642 | -86.2816
6.7645 | -86.2834
6.7648 | -86.2848
6.7651 | -86.2853
6.7654 | -86.2854
6.7657 | -86.2853
6.766 | -86.2852
6.7663 | -86.2852
6.7666 | -86.2853
6.767 | -86.2853
6.7673 | -86.2851
6.7676 | -86.2854
6.7679 | -86.286
6.7682 | -86.286
6.7685 | -86.2861
6.7688 | -86.2868
6.7691 | -86.2874
6.7694 | -86.2874
6.7697 | -86.2876
6.77 | -86.2876
6.7704 | -86.2881
6.7707 | -86.2877
6.771 | -86.2874
6.7713 | -86.2874
6.7716 | -86.2874
6.7719 | -86.2882
6.7722 | -86.2893
6.7725 | -86.29
6.7728 | -86.2903
6.7731 | -86.2892
6.7734 | -86.2872
6.7737 | -86.2858
6.7741 | -86.2854
6.7744 | -86.2852
6.7747 | -86.2853
6.775 | -86.2853
6.7753 | -86.2854
6.7756 | -86.2853
6.7759 | -86.2853
6.7762 | -86.285
6.7765 | -86.2846
6.7768 | -86.2838
6.7771 | -86.2835
6.7775 | -86.2844
6.7778 | -86.2852
6.7781 | -86.2856
6.7784 | -86.2853
6.7787 | -86.2856
6.779 | -86.2867
6.7793 | -86.2871
6.7796 | -86.2873
6.7799 | -86.2875
6.7802 | -86.2876
6.7805 | -86.2876
6.7808 | -86.2876
6.7812 | -86.2882
6.7815 | -86.2893
6.7818 | -86.2888
6.7821 | -86.2873
6.7824 | -86.2867
6.7827 | -86.2864
6.783 | -86.2861
6.7833 | -86.2859
6.7836 | -86.2861
6.7839 | -86.2865
6.7842 | -86.2873
6.7846 | -86.2883
6.7849 | -86.289
6.7852 | -86.2884
6.7855 | -86.2874
6.7858 | -86.2872
6.7861 | -86.2876
6.7864 | -86.2877
6.7867 | -86.2882
6.787 | -86.2882
6.7873 | -86.2874
6.7876 | -86.2874
6.788 | -86.2874
6.7883 | -86.288
6.7886 | -86.2884
6.7889 | -86.2887
6.7892 | -86.2889
6.7895 | -86.2894
6.7898 | -86.2895
6.7901 | -86.2896
6.7904 | -86.2896
6.7907 | -86.289
6.791 | -86.2889
6.7913 | -86.2895
6.7917 | -86.2899
6.792 | -86.2892
6.7923 | -86.2893
6.7926 | -86.289
6.7929 | -86.2888
6.7932 | -86.2888
6.7935 | -86.2887
6.7938 | -86.2886
6.7941 | -86.2885
6.7944 | -86.2887
6.7947 | -86.289
6.7951 | -86.2887
6.7954 | -86.2886
6.7957 | -86.2889
6.796 | -86.2893
6.7963 | -86.2892
6.7966 | -86.2895
6.7969 | -86.2905
6.7972 | -86.2913
6.7975 | -86.292
6.7978 | -86.2923
6.7981 | -86.2916
6.7984 | -86.2918
6.7988 | -86.292
6.7991 | -86.2926
6.7994 | -86.2929
6.7997 | -86.293
6.8 | -86.2927
6.8003 | -86.2926
6.8006 | -86.2923
6.8009 | -86.292
6.8012 | -86.2921
6.8015 | -86.2919
6.8018 | -86.2919
6.8022 | -86.2922
6.8025 | -86.2923
6.8028 | -86.2926
6.8031 | -86.2931
6.8034 | -86.293
6.8037 | -86.2924
6.804 | -86.2926
6.8043 | -86.293
6.8046 | -86.2931
6.8049 | -86.2933
6.8052 | -86.294
6.8056 | -86.2944
6.8059 | -86.294
6.8062 | -86.2943
6.8065 | -86.2936
6.8068 | -86.2937
6.8071 | -86.294
6.8074 | -86.2938
6.8077 | -86.2938
6.808 | -86.2937
6.8083 | -86.294
6.8086 | -86.2952
6.8089 | -86.2973
6.8093 | -86.2981
6.8096 | -86.298
6.8099 | -86.2975
6.8102 | -86.2977
6.8105 | -86.2975
6.8108 | -86.2979
6.8111 | -86.2986
6.8114 | -86.2991
6.8117 | -86.2992
6.812 | -86.2994
6.8123 | -86.2997
6.8127 | -86.3
6.813 | -86.3005
6.8133 | -86.3008
6.8136 | -86.3007
6.8139 | -86.3009
6.8142 | -86.3016
6.8145 | -86.3026
6.8148 | -86.3034
6.8151 | -86.3043
6.8154 | -86.3043
6.8157 | -86.3049
6.816 | -86.3055
6.8164 | -86.3057
6.8167 | -86.3064
6.817 | -86.3076
6.8173 | -86.3087
6.8176 | -86.3092
6.8179 | -86.3095
6.8182 | -86.3107
6.8185 | -86.3113
6.8188 | -86.3114
6.8191 | -86.3117
6.8194 | -86.3122
6.8198 | -86.3132
6.8201 | -86.3144
6.8204 | -86.3152
6.8207 | -86.3144
6.821 | -86.3142
6.8213 | -86.3141
6.8216 | -86.314
6.8219 | -86.314
6.8222 | -86.3143
6.8225 | -86.315
6.8228 | -86.3153
6.8232 | -86.3152
6.8235 | -86.3155
6.8238 | -86.3167
6.8241 | -86.3182
6.8244 | -86.3192
6.8247 | -86.3191
6.825 | -86.3194
6.8253 | -86.3198
6.8256 | -86.3202
6.8259 | -86.3205
6.8262 | -86.321
6.8265 | -86.3214
6.8269 | -86.3216
6.8272 | -86.3217
6.8275 | -86.3222
6.8278 | -86.3225
6.8281 | -86.3221
6.8284 | -86.3222
6.8287 | -86.3224
6.829 | -86.3222
6.8293 | -86.3226
6.8296 | -86.3226
6.8299 | -86.3227
6.8303 | -86.3228
6.8306 | -86.3234
6.8309 | -86.3249
6.8312 | -86.3261
6.8315 | -86.3267
6.8318 | -86.3269
6.8321 | -86.327
6.8324 | -86.327
6.8327 | -86.3281
6.833 | -86.3286
6.8333 | -86.3289
6.8336 | -86.3293
6.834 | -86.3296
6.8343 | -86.3297
6.8346 | -86.3304
6.8349 | -86.3308
6.8352 | -86.3312
6.8355 | -86.3312
6.8358 | -86.331
6.8361 | -86.3318
6.8364 | -86.3322
6.8367 | -86.3326
6.837 | -86.3331
6.8374 | -86.3335
6.8377 | -86.3331
6.838 | -86.3328
6.8383 | -86.3332
6.8386 | -86.3334
6.8389 | -86.3329
6.8392 | -86.3324
6.8395 | -86.3324
6.8398 | -86.3329
6.8401 | -86.3331
6.8404 | -86.3333
6.8407 | -86.3336
6.8411 | -86.3341
6.8414 | -86.3345
6.8417 | -86.3351
6.842 | -86.3349
6.8423 | -86.3353
6.8426 | -86.3357
6.8429 | -86.3359
6.8432 | -86.3365
6.8435 | -86.3367
6.8438 | -86.3371
6.8441 | -86.3375
6.8445 | -86.3383
6.8448 | -86.3388
6.8451 | -86.3396
6.8454 | -86.3405
6.8457 | -86.3414
6.846 | -86.3422
6.8463 | -86.3428
6.8466 | -86.3435
6.8469 | -86.3442
6.8472 | -86.3452
6.8475 | -86.3457
6.8479 | -86.3463
6.8482 | -86.3467
6.8485 | -86.3472
6.8488 | -86.3478
6.8491 | -86.3489
6.8494 | -86.3497
6.8497 | -86.3502
6.85 | -86.3508
6.8503 | -86.3514
6.8506 | -86.3528
6.8509 | -86.353
6.8512 | -86.3536
6.8516 | -86.3554
6.8519 | -86.3559
6.8522 | -86.3566
6.8525 | -86.3578
6.8528 | -86.3588
6.8531 | -86.3596
6.8534 | -86.36
6.8537 | -86.3604
6.854 | -86.3611
6.8543 | -86.3613
6.8546 | -86.3615
6.855 | -86.3625
6.8553 | -86.3638
6.8556 | -86.3644
6.8559 | -86.3654
6.8562 | -86.3656
6.8565 | -86.3659
6.8568 | -86.3665
6.8571 | -86.3672
6.8574 | -86.3678
6.8577 | -86.3683
6.858 | -86.3686
6.8583 | -86.3691
6.8587 | -86.3695
6.859 | -86.3694
6.8593 | -86.3699
6.8596 | -86.3707
6.8599 | -86.371
6.8602 | -86.3717
6.8605 | -86.3726
6.8608 | -86.3742
6.8611 | -86.3748
6.8614 | -86.3752
6.8617 | -86.3753
6.8621 | -86.3759
6.8624 | -86.3763
6.8627 | -86.3768
6.863 | -86.3773
6.8633 | -86.3782
6.8636 | -86.379
6.8639 | -86.3798
6.8642 | -86.3803
6.8645 | -86.3807
6.8648 | -86.3811
6.8651 | -86.3817
6.8655 | -86.3822
6.8658 | -86.3825
6.8661 | -86.3833
6.8664 | -86.3843
6.8667 | -86.3849
6.867 | -86.3856
6.8673 | -86.3861
6.8676 | -86.3865
6.8679 | -86.3872
6.8682 | -86.3876
6.8685 | -86.3887
6.8688 | -86.3892
6.8692 | -86.3901
6.8695 | -86.3909
6.8698 | -86.3917
6.8701 | -86.3926
6.8704 | -86.3932
6.8707 | -86.3939
6.871 | -86.3944
6.8713 | -86.3945
6.8716 | -86.3953
6.8719 | -86.396
6.8722 | -86.3969
6.8726 | -86.398
6.8729 | -86.3989
6.8732 | -86.4001
6.8735 | -86.4012
6.8738 | -86.4025
6.8741 | -86.4047
6.8744 | -86.406
6.8747 | -86.4059
6.875 | -86.4067
6.8753 | -86.4077
6.8756 | -86.4089
6.8759 | -86.4099
6.8763 | -86.4106
6.8766 | -86.4108
6.8769 | -86.4117
6.8772 | -86.4128
6.8775 | -86.4145
6.8778 | -86.4154
6.8781 | -86.4165
6.8784 | -86.4176
6.8787 | -86.4185
6.879 | -86.4195
6.8793 | -86.4205
6.8797 | -86.4218
6.88 | -86.4229
6.8803 | -86.4238
6.8806 | -86.4247
6.8809 | -86.4254
6.8812 | -86.4261
6.8815 | -86.4265
6.8818 | -86.4273
6.8821 | -86.4282
6.8824 | -86.4293
6.8827 | -86.4303
6.8831 | -86.4311
6.8834 | -86.432
6.8837 | -86.4333
6.884 | -86.4347
6.8843 | -86.4357
6.8846 | -86.4373
6.8849 | -86.4393
6.8852 | -86.4406
6.8855 | -86.4417
6.8858 | -86.4425
6.8861 | -86.4442
6.8864 | -86.4455
6.8868 | -86.4472
6.8871 | -86.4486
6.8874 | -86.45
6.8877 | -86.4518
6.888 | -86.4535
6.8883 | -86.4546
6.8886 | -86.4567
6.8889 | -86.4584
6.8892 | -86.4595
6.8895 | -86.4605
6.8898 | -86.4612
6.8902 | -86.4621
6.8905 | -86.4633
6.8908 | -86.4643
6.8911 | -86.4657
6.8914 | -86.467
6.8917 | -86.468
6.892 | -86.4691
6.8923 | -86.4708
6.8926 | -86.4728
6.8929 | -86.4749
6.8932 | -86.4765
6.8935 | -86.4776
6.8939 | -86.4785
6.8942 | -86.4795
6.8945 | -86.4804
6.8948 | -86.4814
6.8951 | -86.4831
6.8954 | -86.4839
6.8957 | -86.4852
6.896 | -86.4859
6.8963 | -86.4869
6.8966 | -86.4878
6.8969 | -86.4888
6.8973 | -86.4899
6.8976 | -86.4913
6.8979 | -86.4927
6.8982 | -86.4937
6.8985 | -86.4945
6.8988 | -86.4955
6.8991 | -86.4966
6.8994 | -86.4977
6.8997 | -86.4991
6.9 | -86.5001
6.9003 | -86.5011
6.9007 | -86.5019
6.901 | -86.5034
6.9013 | -86.5039
6.9016 | -86.5044
6.9019 | -86.5055
6.9022 | -86.507
6.9025 | -86.5083
6.9028 | -86.5097
6.9031 | -86.512
6.9034 | -86.5133
6.9037 | -86.5146
6.904 | -86.5161
6.9044 | -86.5173
6.9047 | -86.5187
6.905 | -86.5206
6.9053 | -86.5219
6.9056 | -86.5235
6.9059 | -86.5254
6.9062 | -86.5271
6.9065 | -86.5287
6.9068 | -86.5292
6.9071 | -86.5301
6.9074 | -86.5314
6.9078 | -86.5329
6.9081 | -86.5352
6.9084 | -86.5369
6.9087 | -86.538
6.909 | -86.5392
6.9093 | -86.5402
6.9096 | -86.5418
6.9099 | -86.5446
6.9102 | -86.547
6.9105 | -86.5482
6.9108 | -86.5495
6.9111 | -86.5509
6.9115 | -86.5523
6.9118 | -86.5533
6.9121 | -86.5539
6.9124 | -86.555
6.9127 | -86.5566
6.913 | -86.5587
6.9133 | -86.5605
6.9136 | -86.5624
6.9139 | -86.5641
6.9142 | -86.5651
6.9145 | -86.5663
6.9149 | -86.5677
6.9152 | -86.569
6.9155 | -86.5703
6.9158 | -86.5716
6.9161 | -86.5728
6.9164 | -86.5743
6.9167 | -86.577
6.917 | -86.5794
6.9173 | -86.5809
6.9176 | -86.5822
6.9179 | -86.5834
6.9183 | -86.5845
6.9186 | -86.5851
6.9189 | -86.5861
6.9192 | -86.587
6.9195 | -86.5878
6.9198 | -86.5886
6.9201 | -86.5894
6.9204 | -86.5909
6.9207 | -86.5922
6.921 | -86.5936
6.9213 | -86.5959
6.9216 | -86.5973
6.922 | -86.5991
6.9223 | -86.6006
6.9226 | -86.6024
6.9229 | -86.6037
6.9232 | -86.605
6.9235 | -86.6064
6.9238 | -86.6077
6.9241 | -86.6091
6.9244 | -86.6105
6.9247 | -86.6119
6.925 | -86.6135
6.9254 | -86.615
6.9257 | -86.6165
6.926 | -86.6177
6.9263 | -86.6185
6.9266 | -86.6193
6.9269 | -86.6205
6.9272 | -86.6222
6.9275 | -86.6235
6.9278 | -86.6251
6.9281 | -86.6269
6.9284 | -86.6282
6.9287 | -86.63
6.9291 | -86.6313
6.9294 | -86.6325
6.9297 | -86.6337
6.93 | -86.6358
6.9303 | -86.6381
6.9306 | -86.64
6.9309 | -86.6411
6.9312 | -86.6426
6.9315 | -86.6438
6.9318 | -86.6451
6.9321 | -86.6456
6.9325 | -86.647
6.9328 | -86.6482
6.9331 | -86.6493
6.9334 | -86.6507
6.9337 | -86.6522
6.934 | -86.6536
6.9343 | -86.6551
6.9346 | -86.6566
6.9349 | -86.6579
6.9352 | -86.6595
6.9355 | -86.6609
6.9359 | -86.6623
6.9362 | -86.6637
6.9365 | -86.665
6.9368 | -86.6663
6.9371 | -86.6678
6.9374 | -86.6691
6.9377 | -86.67
6.938 | -86.6715
6.9383 | -86.6734
6.9386 | -86.6746
6.9389 | -86.676
6.9392 | -86.6776
6.9396 | -86.6792
6.9399 | -86.6813
6.9402 | -86.6831
6.9405 | -86.6851
6.9408 | -86.6865
6.9411 | -86.6882
6.9414 | -86.69
6.9417 | -86.6914
6.942 | -86.6927
6.9423 | -86.6944
6.9426 | -86.6956
6.943 | -86.6971
6.9433 | -86.6987
6.9436 | -86.7002
6.9439 | -86.7021
6.9442 | -86.704
6.9445 | -86.7057
6.9448 | -86.7071
6.9451 | -86.7093
6.9454 | -86.7108
6.9457 | -86.7124
6.946 | -86.714
6.9463 | -86.7152
6.9467 | -86.7168
6.947 | -86.7187
6.9473 | -86.7204
6.9476 | -86.7218
6.9479 | -86.723
6.9482 | -86.7244
6.9485 | -86.7265
6.9488 | -86.7283
6.9491 | -86.7301
6.9494 | -86.7323
6.9497 | -86.7341
6.9501 | -86.7359
6.9504 | -86.7378
6.9507 | -86.7395
6.951 | -86.7411
6.9513 | -86.7423
6.9516 | -86.7439
6.9519 | -86.7458
6.9522 | -86.7477
6.9525 | -86.7495
6.9528 | -86.7511
6.9531 | -86.7528
6.9535 | -86.7545
6.9538 | -86.7558
6.9541 | -86.7575
6.9544 | -86.7593
6.9547 | -86.7609
6.955 | -86.763
6.9553 | -86.765
6.9556 | -86.7667
6.9559 | -86.7682
6.9562 | -86.7696
6.9565 | -86.7714
6.9568 | -86.7733
6.9572 | -86.775
6.9575 | -86.7771
6.9578 | -86.7789
6.9581 | -86.7816
6.9584 | -86.784
6.9587 | -86.7854
6.959 | -86.7866
6.9593 | -86.7882
6.9596 | -86.7897
6.9599 | -86.7916
6.9602 | -86.7933
6.9606 | -86.7955
6.9609 | -86.7983
6.9612 | -86.8005
6.9615 | -86.8022
6.9618 | -86.8037
6.9621 | -86.8059
6.9624 | -86.8081
6.9627 | -86.81
6.963 | -86.8126
6.9633 | -86.8149
6.9636 | -86.8163
6.9639 | -86.8176
6.9643 | -86.8193
6.9646 | -86.8209
6.9649 | -86.8224
6.9652 | -86.8242
6.9655 | -86.8259
6.9658 | -86.8277
6.9661 | -86.8296
6.9664 | -86.832
6.9667 | -86.834
6.967 | -86.8359
6.9673 | -86.8376
6.9677 | -86.8394
6.968 | -86.841
6.9683 | -86.8427
6.9686 | -86.8449
6.9689 | -86.8473
6.9692 | -86.8494
6.9695 | -86.8519
6.9698 | -86.854
6.9701 | -86.8561
6.9704 | -86.8584
6.9707 | -86.8598
6.971 | -86.8616
6.9714 | -86.8632
6.9717 | -86.8652
6.972 | -86.8673
6.9723 | -86.869
6.9726 | -86.8706
6.9729 | -86.8727
6.9732 | -86.8747
6.9735 | -86.8765
6.9738 | -86.8785
6.9741 | -86.8802
6.9744 | -86.8828
6.9748 | -86.8845
6.9751 | -86.8859
6.9754 | -86.8878
6.9757 | -86.8902
6.976 | -86.8936
6.9763 | -86.8974
6.9766 | -86.9004
6.9769 | -86.9017
6.9772 | -86.9031
6.9775 | -86.905
6.9778 | -86.9068
6.9782 | -86.9072
6.9785 | -86.9092
6.9788 | -86.9111
6.9791 | -86.9128
6.9794 | -86.9144
6.9797 | -86.9161
6.98 | -86.9182
6.9803 | -86.9202
6.9806 | -86.922
6.9809 | -86.9234
6.9812 | -86.9249
6.9815 | -86.9254
6.9819 | -86.9258
6.9822 | -86.9268
6.9825 | -86.9285
6.9828 | -86.9303
6.9831 | -86.9321
6.9834 | -86.9341
6.9837 | -86.9366
6.984 | -86.9385
6.9843 | -86.9409
6.9846 | -86.9425
6.9849 | -86.9449
6.9853 | -86.9469
6.9856 | -86.9479
6.9859 | -86.95
6.9862 | -86.9523
6.9865 | -86.9545
6.9868 | -86.9582
6.9871 | -86.9608
6.9874 | -86.9624
6.9877 | -86.964
6.988 | -86.9657
6.9883 | -86.9675
6.9886 | -86.9692
6.989 | -86.9711
6.9893 | -86.974
6.9896 | -86.9761
6.9899 | -86.9779
6.9902 | -86.98
6.9905 | -86.9819
6.9908 | -86.9837
6.9911 | -86.9856
6.9914 | -86.9875
6.9917 | -86.9895
6.992 | -86.9914
6.9924 | -86.9931
6.9927 | -86.9948
6.993 | -86.9963
6.9933 | -86.9986
6.9936 | -87.0006
6.9939 | -87.0026
6.9942 | -87.0047
6.9945 | -87.0071
6.9948 | -87.009
6.9951 | -87.0108
6.9954 | -87.0129
6.9958 | -87.0146
6.9961 | -87.0167
6.9964 | -87.0191
6.9967 | -87.0211
6.997 | -87.0228
6.9973 | -87.0244
6.9976 | -87.0267
6.9979 | -87.0287
6.9982 | -87.0309
6.9985 | -87.0337
6.9988 | -87.0362
6.9991 | -87.038
6.9995 | -87.0396
6.9998 | -87.0418
7.0001 | -87.0438
7.0004 | -87.0456
7.0007 | -87.0476
7.001 | -87.0495
7.0013 | -87.0514
7.0016 | -87.0537
7.0019 | -87.0563
7.0022 | -87.0587
7.0025 | -87.0608
7.0029 | -87.0626
7.0032 | -87.0642
7.0035 | -87.0658
7.0038 | -87.0674
7.0041 | -87.0688
7.0044 | -87.0701
7.0047 | -87.0714
7.005 | -87.0736
7.0053 | -87.0763
7.0056 | -87.0799
7.0059 | -87.0835
7.0062 | -87.0861
7.0066 | -87.0881
7.0069 | -87.09
7.0072 | -87.0917
7.0075 | -87.0934
7.0078 | -87.0955
7.0081 | -87.099
7.0084 | -87.1014
7.0087 | -87.1035
7.009 | -87.105
7.0093 | -87.107
7.0096 | -87.1088
7.01 | -87.1106
7.0103 | -87.1123
7.0106 | -87.1143
7.0109 | -87.1169
7.0112 | -87.1187
7.0115 | -87.1203
7.0118 | -87.1225
7.0121 | -87.1245
7.0124 | -87.1264
7.0127 | -87.1282
7.013 | -87.1299
7.0134 | -87.1321
7.0137 | -87.1343
7.014 | -87.136
7.0143 | -87.1378
7.0146 | -87.1398
7.0149 | -87.1414
7.0152 | -87.1429
7.0155 | -87.1444
7.0158 | -87.1465
7.0161 | -87.1495
7.0164 | -87.1519
7.0167 | -87.1538
7.0171 | -87.156
7.0174 | -87.1589
7.0177 | -87.1612
7.018 | -87.1633
7.0183 | -87.1652
7.0186 | -87.1673
7.0189 | -87.1694
7.0192 | -87.1716
7.0195 | -87.1737
7.0198 | -87.1755
7.0201 | -87.1773
7.0205 | -87.1803
7.0208 | -87.1825
7.0211 | -87.1851
7.0214 | -87.1863
7.0217 | -87.1874
7.022 | -87.1898
7.0223 | -87.1915
7.0226 | -87.1943
7.0229 | -87.1978
7.0232 | -87.2004
7.0235 | -87.2027
7.0238 | -87.2053
7.0242 | -87.2072
7.0245 | -87.209
7.0248 | -87.2107
7.0251 | -87.2125
7.0254 | -87.2141
7.0257 | -87.2158
7.026 | -87.2177
7.0263 | -87.2197
7.0266 | -87.2218
7.0269 | -87.224
7.0272 | -87.2267
7.0276 | -87.2294
7.0279 | -87.2314
7.0282 | -87.2334
7.0285 | -87.2356
7.0288 | -87.2378
7.0291 | -87.2399
7.0294 | -87.2424
7.0297 | -87.2449
7.03 | -87.2472
7.0303 | -87.2492
7.0306 | -87.2506
7.031 | -87.2521
7.0313 | -87.2545
7.0316 | -87.2565
7.0319 | -87.2586
7.0322 | -87.2607
7.0325 | -87.2628
7.0328 | -87.265
7.0331 | -87.2668
7.0334 | -87.2698
7.0337 | -87.2728
7.034 | -87.2749
7.0343 | -87.2778
7.0347 | -87.2802
7.035 | -87.2821
7.0353 | -87.2842
7.0356 | -87.2869
7.0359 | -87.2895
7.0362 | -87.2918
7.0365 | -87.2949
7.0368 | -87.2979
7.0371 | -87.2999
7.0374 | -87.3026
7.0377 | -87.3051
7.0381 | -87.3073
7.0384 | -87.3097
7.0387 | -87.3119
7.039 | -87.3138
7.0393 | -87.3153
7.0396 | -87.3171
7.0399 | -87.3191
7.0402 | -87.3213
7.0405 | -87.3238
7.0408 | -87.3257
7.0411 | -87.3278
7.0414 | -87.3306
7.0418 | -87.3333
7.0421 | -87.3352
7.0424 | -87.3371
7.0427 | -87.3392
7.043 | -87.3414
7.0433 | -87.3436
7.0436 | -87.3458
7.0439 | -87.3478
7.0442 | -87.3494
7.0445 | -87.3515
7.0448 | -87.3536
7.0452 | -87.3563
7.0455 | -87.359
7.0458 | -87.3617
7.0461 | -87.3641
7.0464 | -87.3666
7.0467 | -87.3674
7.047 | -87.3698
7.0473 | -87.3717
7.0476 | -87.3741
7.0479 | -87.3761
7.0482 | -87.3781
7.0486 | -87.3803
7.0489 | -87.3824
7.0492 | -87.3845
7.0495 | -87.3865
7.0498 | -87.3888
7.0501 | -87.3919
7.0504 | -87.3942
7.0507 | -87.3958
7.051 | -87.3984
7.0513 | -87.4008
7.0516 | -87.4027
7.0519 | -87.4049
7.0523 | -87.4071
7.0526 | -87.4091
7.0529 | -87.4114
7.0532 | -87.4144
7.0535 | -87.4168
7.0538 | -87.4191
7.0541 | -87.4217
7.0544 | -87.4246
7.0547 | -87.4269
7.055 | -87.4289
7.0553 | -87.4305
7.0557 | -87.4321
7.056 | -87.4351
7.0563 | -87.4372
7.0566 | -87.4397
7.0569 | -87.4426
7.0572 | -87.4443
7.0575 | -87.4462
7.0578 | -87.4488
7.0581 | -87.4507
7.0584 | -87.4534
7.0587 | -87.4564
7.059 | -87.4589
7.0594 | -87.4607
7.0597 | -87.4624
7.06 | -87.4645
7.0603 | -87.4667
7.0606 | -87.4692
7.0609 | -87.4714
7.0612 | -87.4738
7.0615 | -87.4763
7.0618 | -87.4784
7.0621 | -87.4811
7.0624 | -87.4848
7.0628 | -87.4874
7.0631 | -87.4896
7.0634 | -87.4917
7.0637 | -87.4938
7.064 | -87.4959
7.0643 | -87.4985
7.0646 | -87.5013
7.0649 | -87.5042
7.0652 | -87.506
7.0655 | -87.5077
7.0658 | -87.5093
7.0662 | -87.511
7.0665 | -87.5128
7.0668 | -87.5146
7.0671 | -87.5167
7.0674 | -87.5187
7.0677 | -87.5207
7.068 | -87.5234
7.0683 | -87.5274
7.0686 | -87.5298
7.0689 | -87.532
7.0692 | -87.5347
7.0695 | -87.537
7.0699 | -87.539
7.0702 | -87.541
7.0705 | -87.5433
7.0708 | -87.5454
7.0711 | -87.5474
7.0714 | -87.5497
7.0717 | -87.5518
7.072 | -87.5537
7.0723 | -87.5558
7.0726 | -87.558
7.0729 | -87.5599
7.0733 | -87.5623
7.0736 | -87.5648
7.0739 | -87.5673
7.0742 | -87.5694
7.0745 | -87.5714
7.0748 | -87.5733
7.0751 | -87.5749
7.0754 | -87.5765
7.0757 | -87.5782
7.076 | -87.5803
7.0763 | -87.5824
7.0766 | -87.5842
7.077 | -87.586
7.0773 | -87.5881
7.0776 | -87.5901
7.0779 | -87.5921
7.0782 | -87.5943
7.0785 | -87.5968
7.0788 | -87.6009
7.0791 | -87.6045
7.0794 | -87.6072
7.0797 | -87.6096
7.08 | -87.612
7.0804 | -87.6155
7.0807 | -87.6203
7.081 | -87.6235
7.0813 | -87.6256
7.0816 | -87.6276
7.0819 | -87.6294
7.0822 | -87.6313
7.0825 | -87.6334
7.0828 | -87.6355
7.0831 | -87.6381
7.0834 | -87.6408
7.0838 | -87.6436
7.0841 | -87.6476
7.0844 | -87.6509
7.0847 | -87.6531
7.085 | -87.6555
7.0853 | -87.6578
7.0856 | -87.6597
7.0859 | -87.662
7.0862 | -87.6648
7.0865 | -87.6675
7.0868 | -87.6707
7.0871 | -87.674
7.0875 | -87.6767
7.0878 | -87.6793
7.0881 | -87.6815
7.0884 | -87.6836
7.0887 | -87.6855
7.089 | -87.6876
7.0893 | -87.6898
7.0896 | -87.6921
7.0899 | -87.6946
7.0902 | -87.6966
7.0905 | -87.6988
7.0909 | -87.7007
7.0912 | -87.7031
7.0915 | -87.7052
7.0918 | -87.7077
7.0921 | -87.7099
7.0924 | -87.7118
7.0927 | -87.714
7.093 | -87.7164
7.0933 | -87.719
7.0936 | -87.7211
7.0939 | -87.7234
7.0942 | -87.7253
7.0946 | -87.7262
7.0949 | -87.7283
7.0952 | -87.7303
7.0955 | -87.7312
7.0958 | -87.7325
7.0961 | -87.7345
7.0964 | -87.7367
7.0967 | -87.7388
7.097 | -87.7408
7.0973 | -87.7428
7.0976 | -87.745
7.098 | -87.7472
7.0983 | -87.7492
7.0986 | -87.7513
7.0989 | -87.754
7.0992 | -87.7568
7.0995 | -87.7592
7.0998 | -87.7612
7.1001 | -87.7632
7.1004 | -87.765
7.1007 | -87.7675
7.101 | -87.7697
7.1014 | -87.7714
7.1017 | -87.7734
7.102 | -87.7755
7.1023 | -87.7783
7.1026 | -87.7815
7.1029 | -87.7841
7.1032 | -87.7869
7.1035 | -87.789
7.1038 | -87.7925
7.1041 | -87.7966
7.1044 | -87.7993
7.1047 | -87.8009
7.1051 | -87.8033
7.1054 | -87.8055
7.1057 | -87.8083
7.106 | -87.8111
7.1063 | -87.8135
7.1066 | -87.8156
7.1069 | -87.8177
7.1072 | -87.8198
7.1075 | -87.8223
7.1078 | -87.8244
7.1081 | -87.8265
7.1085 | -87.8286
7.1088 | -87.831
7.1091 | -87.8326
7.1094 | -87.8352
7.1097 | -87.8371
7.11 | -87.8392
7.1103 | -87.8412
7.1106 | -87.8435
7.1109 | -87.8461
7.1112 | -87.8487
7.1115 | -87.8507
7.1118 | -87.8532
7.1122 | -87.8554
7.1125 | -87.8577
7.1128 | -87.8625
7.1131 | -87.8656
7.1134 | -87.8676
7.1137 | -87.8699
7.114 | -87.872
7.1143 | -87.874
7.1146 | -87.8763
7.1149 | -87.879
7.1152 | -87.8809
7.1156 | -87.8829
7.1159 | -87.8853
7.1162 | -87.8876
7.1165 | -87.8896
7.1168 | -87.8918
7.1171 | -87.8947
7.1174 | -87.8979
7.1177 | -87.901
7.118 | -87.9034
7.1183 | -87.9057
7.1186 | -87.9079
7.1189 | -87.9101
7.1193 | -87.9122
7.1196 | -87.9144
7.1199 | -87.9166
7.1202 | -87.9197
7.1205 | -87.9241
7.1208 | -87.9286
7.1211 | -87.9321
7.1214 | -87.9339
7.1217 | -87.9362
7.122 | -87.9394
7.1223 | -87.9415
7.1227 | -87.9435
7.123 | -87.9453
7.1233 | -87.9479
7.1236 | -87.9494
7.1239 | -87.9508
7.1242 | -87.9543
7.1245 | -87.957
7.1248 | -87.9592
7.1251 | -87.9613
7.1254 | -87.9628
7.1257 | -87.9645
7.1261 | -87.9654
7.1264 | -87.9667
7.1267 | -87.9689
7.127 | -87.9734
7.1273 | -87.9773
7.1276 | -87.9794
7.1279 | -87.9814
7.1282 | -87.9837
7.1285 | -87.9852
7.1288 | -87.9864
7.1291 | -87.9883
7.1294 | -87.9903
7.1298 | -87.9926
7.1301 | -87.9957
7.1304 | -87.9972
7.1307 | -87.9992
7.131 | -88.0011
7.1313 | -88.0028
7.1316 | -88.005
7.1319 | -88.0072
7.1322 | -88.0094
7.1325 | -88.0116
7.1328 | -88.0134
7.1332 | -88.0148
7.1335 | -88.0173
7.1338 | -88.0198
7.1341 | -88.0218
7.1344 | -88.0239
7.1347 | -88.0262
7.135 | -88.028
7.1353 | -88.0301
7.1356 | -88.0323
7.1359 | -88.0344
7.1362 | -88.0367
7.1365 | -88.0389
7.1369 | -88.0408
7.1372 | -88.0429
7.1375 | -88.0457
7.1378 | -88.048
7.1381 | -88.0497
7.1384 | -88.0515
7.1387 | -88.0537
7.139 | -88.0564
7.1393 | -88.0592
7.1396 | -88.0625
7.1399 | -88.0657
7.1403 | -88.068
7.1406 | -88.0697
7.1409 | -88.0706
7.1412 | -88.0721
7.1415 | -88.0748
7.1418 | -88.0789
7.1421 | -88.0815
7.1424 | -88.083
7.1427 | -88.0852
7.143 | -88.0871
7.1433 | -88.0893
7.1437 | -88.0922
7.144 | -88.0938
7.1443 | -88.0958
7.1446 | -88.0979
7.1449 | -88.1003
7.1452 | -88.1043
7.1455 | -88.1068
7.1458 | -88.1092
7.1461 | -88.1114
7.1464 | -88.1134
7.1467 | -88.1147
7.147 | -88.1163
7.1474 | -88.1182
7.1477 | -88.1204
7.148 | -88.1226
7.1483 | -88.1248
7.1486 | -88.1273
7.1489 | -88.1296
7.1492 | -88.1315
7.1495 | -88.1334
7.1498 | -88.1352
7.1501 | -88.1372
7.1504 | -88.1393
7.1508 | -88.1414
7.1511 | -88.1437
7.1514 | -88.146
7.1517 | -88.1488
7.152 | -88.1515
7.1523 | -88.1528
7.1526 | -88.1543
7.1529 | -88.1565
7.1532 | -88.1581
7.1535 | -88.1595
7.1538 | -88.1619
7.1541 | -88.1644
7.1545 | -88.1662
7.1548 | -88.168
7.1551 | -88.1706
7.1554 | -88.1747
7.1557 | -88.1775
7.156 | -88.1786
7.1563 | -88.1811
7.1566 | -88.1844
7.1569 | -88.1864
7.1572 | -88.1886
7.1575 | -88.1909
7.1579 | -88.1927
7.1582 | -88.1944
7.1585 | -88.1963
7.1588 | -88.1984
7.1591 | -88.2004
7.1594 | -88.2028
7.1597 | -88.2053
7.16 | -88.2076
7.1603 | -88.2096
7.1606 | -88.2111
7.1609 | -88.213
7.1613 | -88.2149
7.1616 | -88.217
7.1619 | -88.2191
7.1622 | -88.2206
7.1625 | -88.2218
7.1628 | -88.2235
7.1631 | -88.2268
7.1634 | -88.2303
7.1637 | -88.2321
7.164 | -88.2343
7.1643 | -88.236
7.1646 | -88.2381
7.165 | -88.2404
7.1653 | -88.2425
7.1656 | -88.2441
7.1659 | -88.2462
7.1662 | -88.2481
7.1665 | -88.25
7.1668 | -88.2524
7.1671 | -88.2546
7.1674 | -88.2565
7.1677 | -88.2583
7.168 | -88.2609
7.1684 | -88.2638
7.1687 | -88.266
7.169 | -88.2679
7.1693 | -88.2697
7.1696 | -88.2717
7.1699 | -88.2746
7.1702 | -88.2766
7.1705 | -88.2785
7.1708 | -88.2819
7.1711 | -88.2842
7.1714 | -88.2854
7.1717 | -88.287
7.1721 | -88.2886
7.1724 | -88.29
7.1727 | -88.2917
7.173 | -88.2934
7.1733 | -88.2953
7.1736 | -88.2971
7.1739 | -88.2991
7.1742 | -88.301
7.1745 | -88.303
7.1748 | -88.3052
7.1751 | -88.307
7.1755 | -88.3086
7.1758 | -88.3103
7.1761 | -88.3121
7.1764 | -88.3141
7.1767 | -88.3158
7.177 | -88.3192
7.1773 | -88.323
7.1776 | -88.3253
7.1779 | -88.3268
7.1782 | -88.3282
7.1785 | -88.3302
7.1789 | -88.3321
7.1792 | -88.3339
7.1795 | -88.3363
7.1798 | -88.3383
7.1801 | -88.3393
7.1804 | -88.3407
7.1807 | -88.3427
7.181 | -88.3446
7.1813 | -88.3459
7.1816 | -88.3476
7.1819 | -88.35
7.1822 | -88.3526
7.1826 | -88.3546
7.1829 | -88.3567
7.1832 | -88.3581
7.1835 | -88.3599
7.1838 | -88.3625
7.1841 | -88.3642
7.1844 | -88.365
7.1847 | -88.3661
7.185 | -88.3679
7.1853 | -88.3695
7.1856 | -88.3707
7.186 | -88.373
7.1863 | -88.3758
7.1866 | -88.378
7.1869 | -88.3798
7.1872 | -88.3812
7.1875 | -88.3834
7.1878 | -88.3851
7.1881 | -88.3871
7.1884 | -88.3889
7.1887 | -88.3908
7.189 | -88.3922
7.1893 | -88.3937
7.1897 | -88.3961
7.19 | -88.3979
7.1903 | -88.399
7.1906 | -88.4006
7.1909 | -88.4024
7.1912 | -88.4041
7.1915 | -88.4054
7.1918 | -88.4063
7.1921 | -88.4081
7.1924 | -88.4095
7.1927 | -88.4113
7.1931 | -88.4123
7.1934 | -88.4139
7.1937 | -88.4158
7.194 | -88.4179
7.1943 | -88.419
7.1946 | -88.4205
7.1949 | -88.4219
7.1952 | -88.4238
7.1955 | -88.4259
7.1958 | -88.4273
7.1961 | -88.429
7.1964 | -88.4306
7.1968 | -88.4327
7.1971 | -88.4344
7.1974 | -88.4364
7.1977 | -88.4386
7.198 | -88.4402
7.1983 | -88.4425
7.1986 | -88.4454
7.1989 | -88.4469
7.1992 | -88.4487
7.1995 | -88.4506
7.1998 | -88.4523
7.2002 | -88.4542
7.2005 | -88.4558
7.2008 | -88.4576
7.2011 | -88.4594
7.2014 | -88.461
7.2017 | -88.4625
7.202 | -88.4637
7.2023 | -88.4653
7.2026 | -88.467
7.2029 | -88.4701
7.2032 | -88.473
7.2036 | -88.475
7.2039 | -88.4768
7.2042 | -88.4788
7.2045 | -88.4804
7.2048 | -88.4825
7.2051 | -88.4848
7.2054 | -88.4867
7.2057 | -88.4882
7.206 | -88.4899
7.2063 | -88.4917
7.2066 | -88.494
7.2069 | -88.4956
7.2073 | -88.4978
7.2076 | -88.4991
7.2079 | -88.501
7.2082 | -88.5027
7.2085 | -88.5044
7.2088 | -88.5068
7.2091 | -88.5088
7.2094 | -88.5104
7.2097 | -88.5125
7.21 | -88.5146
7.2103 | -88.5162
7.2107 | -88.5172
7.211 | -88.5186
7.2113 | -88.5202
7.2116 | -88.5215
7.2119 | -88.5231
7.2122 | -88.5248
7.2125 | -88.526
7.2128 | -88.5261
7.2131 | -88.5272
7.2134 | -88.5296
7.2137 | -88.5309
7.2141 | -88.5324
7.2144 | -88.5341
7.2147 | -88.5356
7.215 | -88.5371
7.2153 | -88.5381
7.2156 | -88.5406
7.2159 | -88.5417
7.2162 | -88.5442
7.2165 | -88.546
7.2168 | -88.5474
7.2171 | -88.5487
7.2174 | -88.55
7.2178 | -88.5513
7.2181 | -88.5532
7.2184 | -88.5553
7.2187 | -88.5567
7.219 | -88.5576
7.2193 | -88.5594
7.2196 | -88.5612
7.2199 | -88.5626
7.2202 | -88.5641
7.2205 | -88.5664
7.2208 | -88.5683
7.2212 | -88.5697
7.2215 | -88.5712
7.2218 | -88.5726
7.2221 | -88.5741
7.2224 | -88.5761
7.2227 | -88.5784
7.223 | -88.58
7.2233 | -88.5818
7.2236 | -88.5837
7.2239 | -88.5848
7.2242 | -88.5859
7.2245 | -88.5871
7.2249 | -88.5882
7.2252 | -88.5895
7.2255 | -88.5914
7.2258 | -88.594
7.2261 | -88.5963
7.2264 | -88.598
7.2267 | -88.5987
7.227 | -88.5999
7.2273 | -88.6015
7.2276 | -88.603
7.2279 | -88.6046
7.2283 | -88.606
7.2286 | -88.6072
7.2289 | -88.6085
7.2292 | -88.6098
7.2295 | -88.6115
7.2298 | -88.6131
7.2301 | -88.6143
7.2304 | -88.6151
7.2307 | -88.616
7.231 | -88.6173
7.2313 | -88.6185
7.2317 | -88.6191
7.232 | -88.6205
7.2323 | -88.6221
7.2326 | -88.6231
7.2329 | -88.6245
7.2332 | -88.6258
7.2335 | -88.6269
7.2338 | -88.6282
7.2341 | -88.6296
7.2344 | -88.6315
7.2347 | -88.633
7.235 | -88.6335
7.2354 | -88.6344
7.2357 | -88.6359
7.236 | -88.6372
7.2363 | -88.6387
7.2366 | -88.6401
7.2369 | -88.6416
7.2372 | -88.6439
7.2375 | -88.6452
7.2378 | -88.6459
7.2381 | -88.6469
7.2384 | -88.6482
7.2388 | -88.6505
7.2391 | -88.6529
7.2394 | -88.6538
7.2397 | -88.654
7.24 | -88.655
7.2403 | -88.6564
7.2406 | -88.6577
7.2409 | -88.659
7.2412 | -88.6602
7.2415 | -88.6607
7.2418 | -88.6617
7.2421 | -88.6626
7.2425 | -88.6648
7.2428 | -88.6658
7.2431 | -88.6682
7.2434 | -88.6698
7.2437 | -88.6711
7.244 | -88.6723
7.2443 | -88.6734
7.2446 | -88.6749
7.2449 | -88.6763
7.2452 | -88.6772
7.2455 | -88.6781
7.2459 | -88.6795
7.2462 | -88.6807
7.2465 | -88.6819
7.2468 | -88.6836
7.2471 | -88.6872
7.2474 | -88.6904
7.2477 | -88.692
7.248 | -88.6932
7.2483 | -88.6949
7.2486 | -88.6964
7.2489 | -88.6976
7.2492 | -88.6991
7.2496 | -88.6999
7.2499 | -88.7013
7.2502 | -88.7024
7.2505 | -88.7031
7.2508 | -88.7048
7.2511 | -88.7061
7.2514 | -88.7068
7.2517 | -88.7097
7.252 | -88.7117
7.2523 | -88.7126
7.2526 | -88.7137
7.253 | -88.715
7.2533 | -88.7157
7.2536 | -88.7168
7.2539 | -88.7182
7.2542 | -88.7196
7.2545 | -88.7208
7.2548 | -88.7221
7.2551 | -88.7233
7.2554 | -88.7242
7.2557 | -88.7252
7.256 | -88.7264
7.2564 | -88.7278
7.2567 | -88.7293
7.257 | -88.7304
7.2573 | -88.7311
7.2576 | -88.7326
7.2579 | -88.7345
7.2582 | -88.7354
7.2585 | -88.7359
7.2588 | -88.7365
7.2591 | -88.737
7.2594 | -88.7383
7.2597 | -88.7401
7.2601 | -88.7432
7.2604 | -88.7439
7.2607 | -88.7482
7.261 | -88.7514
7.2613 | -88.7525
7.2616 | -88.754
7.2619 | -88.755
7.2622 | -88.756
7.2625 | -88.7571
7.2628 | -88.7582
7.2631 | -88.7594
7.2635 | -88.7608
7.2638 | -88.7626
7.2641 | -88.7648
7.2644 | -88.7669
7.2647 | -88.7679
7.265 | -88.7683
7.2653 | -88.7692
7.2656 | -88.7708
7.2659 | -88.7718
7.2662 | -88.7729
7.2665 | -88.774
7.2668 | -88.7752
7.2672 | -88.7758
7.2675 | -88.7763
7.2678 | -88.7771
7.2681 | -88.778
7.2684 | -88.7791
7.2687 | -88.7801
7.269 | -88.7812
7.2693 | -88.7822
7.2696 | -88.7851
7.2699 | -88.7884
7.2702 | -88.7902
7.2706 | -88.7916
7.2709 | -88.7927
7.2712 | -88.7945
7.2715 | -88.7962
7.2718 | -88.7974
7.2721 | -88.7985
7.2724 | -88.7993
7.2727 | -88.8003
7.273 | -88.8024
7.2733 | -88.8053
7.2736 | -88.8066
7.274 | -88.8076
7.2743 | -88.8085
7.2746 | -88.8095
7.2749 | -88.8104
7.2752 | -88.8127
7.2755 | -88.8138
7.2758 | -88.8151
7.2761 | -88.8161
7.2764 | -88.8171
7.2767 | -88.8184
7.277 | -88.8195
7.2773 | -88.8209
7.2777 | -88.8216
7.278 | -88.8226
7.2783 | -88.8237
7.2786 | -88.8247
7.2789 | -88.8257
7.2792 | -88.8277
7.2795 | -88.829
7.2798 | -88.8299
7.2801 | -88.831
7.2804 | -88.8321
7.2807 | -88.8327
7.2811 | -88.8342
7.2814 | -88.8353
7.2817 | -88.836
7.282 | -88.8366
7.2823 | -88.8373
7.2826 | -88.8382
7.2829 | -88.8395
7.2832 | -88.8411
7.2835 | -88.8428
7.2838 | -88.8432
7.2841 | -88.8441
7.2844 | -88.8452
7.2848 | -88.8464
7.2851 | -88.8472
7.2854 | -88.8482
7.2857 | -88.8486
7.286 | -88.8492
7.2863 | -88.85
7.2866 | -88.8511
7.2869 | -88.8526
7.2872 | -88.8523
7.2875 | -88.8532
7.2878 | -88.854
7.2882 | -88.8544
7.2885 | -88.8551
7.2888 | -88.8561
7.2891 | -88.8575
7.2894 | -88.8585
7.2897 | -88.8597
7.29 | -88.8608
7.2903 | -88.8614
7.2906 | -88.8624
7.2909 | -88.8633
7.2912 | -88.8638
7.2916 | -88.8642
7.2919 | -88.8649
7.2922 | -88.8655
7.2925 | -88.8665
7.2928 | -88.8687
7.2931 | -88.8703
7.2934 | -88.8712
7.2937 | -88.8719
7.294 | -88.8725
7.2943 | -88.8735
7.2946 | -88.8749
7.2949 | -88.8769
7.2953 | -88.8779
7.2956 | -88.8794
7.2959 | -88.8813
7.2962 | -88.8809
7.2965 | -88.8814
7.2968 | -88.8825
7.2971 | -88.8835
7.2974 | -88.8842
7.2977 | -88.885
7.298 | -88.8858
7.2983 | -88.8862
7.2987 | -88.8868
7.299 | -88.8875
7.2993 | -88.8884
7.2996 | -88.8897
7.2999 | -88.891
7.3002 | -88.8919
7.3005 | -88.8925
7.3008 | -88.8931
7.3011 | -88.894
7.3014 | -88.8948
7.3017 | -88.8944
7.302 | -88.8951
7.3024 | -88.8952
7.3027 | -88.8957
7.303 | -88.8964
7.3033 | -88.8977
7.3036 | -88.8987
7.3039 | -88.8993
7.3042 | -88.9
7.3045 | -88.9006
7.3048 | -88.901
7.3051 | -88.9013
7.3054 | -88.9022
7.3058 | -88.9029
7.3061 | -88.9037
7.3064 | -88.9044
7.3067 | -88.9052
7.307 | -88.9064
7.3073 | -88.9074
7.3076 | -88.9078
7.3079 | -88.9081
7.3082 | -88.9092
7.3085 | -88.9097
7.3088 | -88.9101
7.3092 | -88.9099
7.3095 | -88.9105
7.3098 | -88.9111
7.3101 | -88.911
7.3104 | -88.9113
7.3107 | -88.9127
7.311 | -88.9151
7.3113 | -88.916
7.3116 | -88.9169
7.3119 | -88.9185
7.3122 | -88.9196
7.3125 | -88.9206
7.3129 | -88.9216
7.3132 | -88.9224
7.3135 | -88.9233
7.3138 | -88.9237
7.3141 | -88.9242
7.3144 | -88.9248
7.3147 | -88.9253
7.315 | -88.926
7.3153 | -88.927
7.3156 | -88.9279
7.3159 | -88.9285
7.3163 | -88.929
7.3166 | -88.9297
7.3169 | -88.9305
7.3172 | -88.9311
7.3175 | -88.9319
7.3178 | -88.9323
7.3181 | -88.9329
7.3184 | -88.9336
7.3187 | -88.935
7.319 | -88.9369
7.3193 | -88.9379
7.3196 | -88.9389
7.32 | -88.9406
7.3203 | -88.9424
7.3206 | -88.9436
7.3209 | -88.9439
7.3212 | -88.9444
7.3215 | -88.9441
7.3218 | -88.9441
7.3221 | -88.9448
7.3224 | -88.9452
7.3227 | -88.9461
7.323 | -88.9468
7.3234 | -88.9478
7.3237 | -88.9488
7.324 | -88.9495
7.3243 | -88.9502
7.3246 | -88.9509
7.3249 | -88.9514
7.3252 | -88.9517
7.3255 | -88.9523
7.3258 | -88.9532
7.3261 | -88.954
7.3264 | -88.9547
7.3267 | -88.9549
7.3271 | -88.9559
7.3274 | -88.9576
7.3277 | -88.9582
7.328 | -88.9588
7.3283 | -88.9594
7.3286 | -88.9601
7.3289 | -88.9609
7.3292 | -88.963
7.3295 | -88.9647
7.3298 | -88.965
7.3301 | -88.9658
7.3305 | -88.9681
7.3308 | -88.9703
7.3311 | -88.9707
7.3314 | -88.9709
7.3317 | -88.9718
7.332 | -88.9729
7.3323 | -88.9743
7.3326 | -88.9751
7.3329 | -88.9756
7.3332 | -88.9762
7.3335 | -88.9772
7.3339 | -88.9782
7.3342 | -88.9801
7.3345 | -88.9807
7.3348 | -88.9814
7.3351 | -88.9825
7.3354 | -88.983
7.3357 | -88.9841
7.336 | -88.985
7.3363 | -88.9856
7.3366 | -88.9864
7.3369 | -88.987
7.3372 | -88.9876
7.3376 | -88.9885
7.3379 | -88.9888
7.3382 | -88.9893
7.3385 | -88.9898
7.3388 | -88.99
7.3391 | -88.9898
7.3394 | -88.9901
7.3397 | -88.9906
7.34 | -88.9911
7.3403 | -88.9919
7.3406 | -88.9929
7.341 | -88.9937
7.3413 | -88.9947
7.3416 | -88.9953
7.3419 | -88.9957
7.3422 | -88.9962
7.3425 | -88.9968
7.3428 | -88.9975
7.3431 | -88.998
7.3434 | -88.9983
7.3437 | -88.9991
7.344 | -89.0005
7.3444 | -89.001
7.3447 | -89.0016
7.345 | -89.0018
7.3453 | -89.0025
7.3456 | -89.0038
7.3459 | -89.0043
7.3462 | -89.0052
7.3465 | -89.0058
7.3468 | -89.0058
7.3471 | -89.0067
7.3474 | -89.0074
7.3477 | -89.0079
7.3481 | -89.0085
7.3484 | -89.0094
7.3487 | -89.0101
7.349 | -89.0113
7.3493 | -89.0124
7.3496 | -89.0131
7.3499 | -89.0131
7.3502 | -89.0132
7.3505 | -89.0136
7.3508 | -89.0145
7.3511 | -89.0154
7.3515 | -89.0161
7.3518 | -89.0172
7.3521 | -89.0194
7.3524 | -89.0201
7.3527 | -89.021
7.353 | -89.0223
7.3533 | -89.0238
7.3536 | -89.0242
7.3539 | -89.0248
7.3542 | -89.0253
7.3545 | -89.0262
7.3548 | -89.0274
7.3552 | -89.0293
7.3555 | -89.0314
7.3558 | -89.0324
7.3561 | -89.0331
7.3564 | -89.0336
7.3567 | -89.0346
7.357 | -89.035
7.3573 | -89.0354
7.3576 | -89.0362
7.3579 | -89.037
7.3582 | -89.0378
7.3586 | -89.0389
7.3589 | -89.0396
7.3592 | -89.0404
7.3595 | -89.0416
7.3598 | -89.0422
7.3601 | -89.0427
7.3604 | -89.043
7.3607 | -89.0434
7.361 | -89.0439
7.3613 | -89.0448
7.3616 | -89.0457
7.362 | -89.0471
7.3623 | -89.048
7.3626 | -89.0487
7.3629 | -89.0496
7.3632 | -89.05
7.3635 | -89.0502
7.3638 | -89.05
7.3641 | -89.0501
7.3644 | -89.0509
7.3647 | -89.0514
7.365 | -89.0525
7.3653 | -89.0542
7.3657 | -89.0547
7.366 | -89.0553
7.3663 | -89.0557
7.3666 | -89.0563
7.3669 | -89.0566
7.3672 | -89.0576
7.3675 | -89.0579
7.3678 | -89.0589
7.3681 | -89.0599
7.3684 | -89.06
7.3687 | -89.0614
7.3691 | -89.0626
7.3694 | -89.0628
7.3697 | -89.0633
7.37 | -89.0638
7.3703 | -89.0645
7.3706 | -89.0647
7.3709 | -89.066
7.3712 | -89.0666
7.3715 | -89.0671
7.3718 | -89.067
7.3721 | -89.0669
7.3724 | -89.0678
7.3728 | -89.0688
7.3731 | -89.0696
7.3734 | -89.0704
7.3737 | -89.0712
7.374 | -89.0716
7.3743 | -89.072
7.3746 | -89.0723
7.3749 | -89.073
7.3752 | -89.0731
7.3755 | -89.0732
7.3758 | -89.0741
7.3762 | -89.075
7.3765 | -89.0763
7.3768 | -89.0773
7.3771 | -89.0778
7.3774 | -89.0783
7.3777 | -89.0788
7.378 | -89.0792
7.3783 | -89.0796
7.3786 | -89.0802
7.3789 | -89.0807
7.3792 | -89.0814
7.3796 | -89.0819
7.3799 | -89.0824
7.3802 | -89.0832
7.3805 | -89.0838
7.3808 | -89.0844
7.3811 | -89.0849
7.3814 | -89.085
7.3817 | -89.0853
7.382 | -89.0862
7.3823 | -89.087
7.3826 | -89.0876
7.3829 | -89.0882
7.3833 | -89.0884
7.3836 | -89.0886
7.3839 | -89.0891
7.3842 | -89.0903
7.3845 | -89.0912
7.3848 | -89.0917
7.3851 | -89.0928
7.3854 | -89.0937
7.3857 | -89.0948
7.386 | -89.0954
7.3863 | -89.0956
7.3867 | -89.0961
7.387 | -89.0969
7.3873 | -89.0978
7.3876 | -89.0991
7.3879 | -89.0998
7.3882 | -89.1001
7.3885 | -89.1007
7.3888 | -89.1012
7.3891 | -89.1018
7.3894 | -89.1028
7.3897 | -89.1036
7.39 | -89.1046
7.3904 | -89.1051
7.3907 | -89.1057
7.391 | -89.1068
7.3913 | -89.1078
7.3916 | -89.108
7.3919 | -89.1082
7.3922 | -89.1087
7.3925 | -89.1094
7.3928 | -89.11
7.3931 | -89.1106
7.3934 | -89.1105
7.3938 | -89.1108
7.3941 | -89.1117
7.3944 | -89.1124
7.3947 | -89.1135
7.395 | -89.1142
7.3953 | -89.115
7.3956 | -89.1158
7.3959 | -89.1163
7.3962 | -89.1164
7.3965 | -89.117
7.3968 | -89.1186
7.3971 | -89.1193
7.3975 | -89.1212
7.3978 | -89.1229
7.3981 | -89.1239
7.3984 | -89.1245
7.3987 | -89.1251
7.399 | -89.1259
7.3993 | -89.1264
7.3996 | -89.1268
7.3999 | -89.1276
7.4002 | -89.1288
7.4005 | -89.1297
7.4009 | -89.1302
7.4012 | -89.1311
7.4015 | -89.1319
7.4018 | -89.1322
7.4021 | -89.1322
7.4024 | -89.1321
7.4027 | -89.1325
7.403 | -89.1325
7.4033 | -89.1327
7.4036 | -89.1345
7.4039 | -89.1359
7.4043 | -89.1367
7.4046 | -89.1376
7.4049 | -89.1384
7.4052 | -89.1391
7.4055 | -89.1399
7.4058 | -89.1408
7.4061 | -89.1417
7.4064 | -89.1424
7.4067 | -89.1432
7.407 | -89.1437
7.4073 | -89.1445
7.4076 | -89.1448
7.408 | -89.1452
7.4083 | -89.1466
7.4086 | -89.1485
7.4089 | -89.1501
7.4092 | -89.151
7.4095 | -89.1516
7.4098 | -89.1522
7.4101 | -89.1526
7.4104 | -89.1528
7.4107 | -89.1536
7.411 | -89.1551
7.4114 | -89.1566
7.4117 | -89.1577
7.412 | -89.1591
7.4123 | -89.1602
7.4126 | -89.1611
7.4129 | -89.1619
7.4132 | -89.1627
7.4135 | -89.1636
7.4138 | -89.1645
7.4141 | -89.1652
7.4144 | -89.1658
7.4147 | -89.1662
7.4151 | -89.1669
7.4154 | -89.1678
7.4157 | -89.1684
7.416 | -89.1695
7.4163 | -89.1706
7.4166 | -89.1718
7.4169 | -89.1732
7.4172 | -89.175
7.4175 | -89.1765
7.4178 | -89.177
7.4181 | -89.1778
7.4185 | -89.1783
7.4188 | -89.179
7.4191 | -89.1796
7.4194 | -89.1804
7.4197 | -89.1816
7.42 | -89.1831
7.4203 | -89.1842
7.4206 | -89.1852
7.4209 | -89.1856
7.4212 | -89.1866
7.4215 | -89.1874
7.4219 | -89.1883
7.4222 | -89.189
7.4225 | -89.1898
7.4228 | -89.1906
7.4231 | -89.1913
7.4234 | -89.1916
7.4237 | -89.1923
7.424 | -89.1935
7.4243 | -89.1952
7.4246 | -89.1973
7.4249 | -89.1987
7.4252 | -89.1992
7.4256 | -89.2003
7.4259 | -89.2013
7.4262 | -89.202
7.4265 | -89.2028
7.4268 | -89.2035
7.4271 | -89.2045
7.4274 | -89.2056
7.4277 | -89.2065
7.428 | -89.2072
7.4283 | -89.2079
7.4286 | -89.2091
7.429 | -89.2111
7.4293 | -89.2126
7.4296 | -89.2124
7.4299 | -89.2131
7.4302 | -89.2139
7.4305 | -89.2156
7.4308 | -89.218
7.4311 | -89.2189
7.4314 | -89.22
7.4317 | -89.2215
7.432 | -89.2221
7.4323 | -89.2227
7.4327 | -89.2238
7.433 | -89.2247
7.4333 | -89.2255
7.4336 | -89.2262
7.4339 | -89.2266
7.4342 | -89.2271
7.4345 | -89.2275
7.4348 | -89.2272
7.4351 | -89.2279
7.4354 | -89.229
7.4357 | -89.2297
7.4361 | -89.2308
7.4364 | -89.2323
7.4367 | -89.2332
7.437 | -89.2342
7.4373 | -89.2355
7.4376 | -89.2364
7.4379 | -89.2371
7.4382 | -89.2378
7.4385 | -89.2391
7.4388 | -89.2399
7.4391 | -89.2411
7.4395 | -89.2421
7.4398 | -89.2428
7.4401 | -89.2439
7.4404 | -89.2451
7.4407 | -89.2467
7.441 | -89.2479
7.4413 | -89.2493
7.4416 | -89.2505
7.4419 | -89.251
7.4422 | -89.2512
7.4425 | -89.2518
7.4428 | -89.2526
7.4432 | -89.2534
7.4435 | -89.2545
7.4438 | -89.2554
7.4441 | -89.2563
7.4444 | -89.257
7.4447 | -89.2578
7.445 | -89.2587
7.4453 | -89.2597
7.4456 | -89.2609
7.4459 | -89.2617
7.4462 | -89.2617
7.4466 | -89.2622
7.4469 | -89.2633
7.4472 | -89.264
7.4475 | -89.2641
7.4478 | -89.2653
7.4481 | -89.2662
7.4484 | -89.2663
7.4487 | -89.2669
7.449 | -89.268
7.4493 | -89.2692
7.4496 | -89.2696
7.4499 | -89.2696
7.4503 | -89.27
7.4506 | -89.2711
7.4509 | -89.2722
7.4512 | -89.2741
7.4515 | -89.2745
7.4518 | -89.2753
7.4521 | -89.2764
7.4524 | -89.2772
7.4527 | -89.278
7.453 | -89.2786
7.4533 | -89.2786
7.4537 | -89.2796
7.454 | -89.2808
7.4543 | -89.2819
7.4546 | -89.2834
7.4549 | -89.2851
7.4552 | -89.286
7.4555 | -89.2866
7.4558 | -89.2881
7.4561 | -89.2896
7.4564 | -89.2893
7.4567 | -89.2899
7.457 | -89.291
7.4574 | -89.2919
7.4577 | -89.2931
7.458 | -89.2938
7.4583 | -89.2944
7.4586 | -89.2951
7.4589 | -89.296
7.4592 | -89.2969
7.4595 | -89.2983
7.4598 | -89.2993
7.4601 | -89.3004
7.4604 | -89.3014
7.4608 | -89.3023
7.4611 | -89.3034
7.4614 | -89.3045
7.4617 | -89.3055
7.462 | -89.3064
7.4623 | -89.3075
7.4626 | -89.3085
7.4629 | -89.3093
7.4632 | -89.3103
7.4635 | -89.3114
7.4638 | -89.312
7.4642 | -89.3136
7.4645 | -89.3151
7.4648 | -89.3149
7.4651 | -89.3153
7.4654 | -89.3167
7.4657 | -89.318
7.466 | -89.319
7.4663 | -89.3198
7.4666 | -89.3204
7.4669 | -89.3208
7.4672 | -89.3216
7.4675 | -89.3224
7.4679 | -89.3232
7.4682 | -89.3238
7.4685 | -89.3248
7.4688 | -89.3258
7.4691 | -89.3268
7.4694 | -89.3279
7.4697 | -89.3287
7.47 | -89.3287
7.4703 | -89.3292
7.4706 | -89.3302
7.4709 | -89.3313
7.4713 | -89.3319
7.4716 | -89.3327
7.4719 | -89.3342
7.4722 | -89.3351
7.4725 | -89.3359
7.4728 | -89.3368
7.4731 | -89.3378
7.4734 | -89.3386
7.4737 | -89.3388
7.474 | -89.3395
7.4743 | -89.341
7.4746 | -89.3419
7.475 | -89.3429
7.4753 | -89.3437
7.4756 | -89.3447
7.4759 | -89.3456
7.4762 | -89.3473
7.4765 | -89.3492
7.4768 | -89.3504
7.4771 | -89.3519
7.4774 | -89.3532
7.4777 | -89.3541
7.478 | -89.3543
7.4784 | -89.3551
7.4787 | -89.3567
7.479 | -89.358
7.4793 | -89.359
7.4796 | -89.3603
7.4799 | -89.3634
7.4802 | -89.3645
7.4805 | -89.3654
7.4808 | -89.3668
7.4811 | -89.368
7.4814 | -89.3692
7.4818 | -89.3701
7.4821 | -89.3712
7.4824 | -89.3724
7.4827 | -89.3739
7.483 | -89.3754
7.4833 | -89.3768
7.4836 | -89.3793
7.4839 | -89.3828
7.4842 | -89.3847
7.4845 | -89.3858
7.4848 | -89.3868
7.4851 | -89.3878
7.4855 | -89.3886
7.4858 | -89.3896
7.4861 | -89.3906
7.4864 | -89.3915
7.4867 | -89.3925
7.487 | -89.3934
7.4873 | -89.3933
7.4876 | -89.3943
7.4879 | -89.3966
7.4882 | -89.3973
7.4885 | -89.3979
7.4889 | -89.3992
7.4892 | -89.4002
7.4895 | -89.4013
7.4898 | -89.4026
7.4901 | -89.4037
7.4904 | -89.4047
7.4907 | -89.4053
7.491 | -89.4063
7.4913 | -89.4074
7.4916 | -89.4084
7.4919 | -89.4098
7.4923 | -89.4103
7.4926 | -89.4112
7.4929 | -89.4121
7.4932 | -89.4131
7.4935 | -89.4141
7.4938 | -89.4151
7.4941 | -89.416
7.4944 | -89.4168
7.4947 | -89.4182
7.495 | -89.4199
7.4953 | -89.4212
7.4956 | -89.4226
7.496 | -89.4235
7.4963 | -89.4242
7.4966 | -89.4255
7.4969 | -89.4262
7.4972 | -89.427
7.4975 | -89.4276
7.4978 | -89.4287
7.4981 | -89.4298
7.4984 | -89.4308
7.4987 | -89.4322
7.499 | -89.4335
7.4994 | -89.4351
7.4997 | -89.4364
7.5 | -89.4377
7.5003 | -89.439
7.5006 | -89.4401
7.5009 | -89.4416
7.5012 | -89.4427
7.5015 | -89.444
7.5018 | -89.445
7.5021 | -89.4461
7.5024 | -89.4473
7.5027 | -89.4486
7.5031 | -89.4497
7.5034 | -89.4508
7.5037 | -89.4521
7.504 | -89.4531
7.5043 | -89.4542
7.5046 | -89.4553
7.5049 | -89.4563
7.5052 | -89.4574
7.5055 | -89.4587
7.5058 | -89.4593
7.5061 | -89.4601
7.5065 | -89.4615
7.5068 | -89.4631
7.5071 | -89.4647
7.5074 | -89.4664
7.5077 | -89.4683
7.508 | -89.4698
7.5083 | -89.4719
7.5086 | -89.4732
7.5089 | -89.4739
7.5092 | -89.4749
7.5095 | -89.476
7.5099 | -89.4777
7.5102 | -89.4786
7.5105 | -89.4797
7.5108 | -89.4808
7.5111 | -89.482
7.5114 | -89.4835
7.5117 | -89.4848
7.512 | -89.4855
7.5123 | -89.486
7.5126 | -89.4869
7.5129 | -89.4884
7.5132 | -89.4894
7.5136 | -89.4901
7.5139 | -89.4916
7.5142 | -89.4934
7.5145 | -89.4943
7.5148 | -89.4954
7.5151 | -89.4963
7.5154 | -89.4973
7.5157 | -89.499
7.516 | -89.4999
7.5163 | -89.5009
7.5166 | -89.5018
7.517 | -89.5029
7.5173 | -89.5043
7.5176 | -89.5047
7.5179 | -89.5052
7.5182 | -89.5066
7.5185 | -89.5077
7.5188 | -89.5093
7.5191 | -89.5099
7.5194 | -89.5106
7.5197 | -89.5119
7.52 | -89.5143
7.5203 | -89.5163
7.5207 | -89.5172
7.521 | -89.5181
7.5213 | -89.5196
7.5216 | -89.5211
7.5219 | -89.522
7.5222 | -89.5229
7.5225 | -89.5235
7.5228 | -89.5237
7.5231 | -89.5246
7.5234 | -89.5253
7.5237 | -89.5261
7.5241 | -89.5273
7.5244 | -89.5289
7.5247 | -89.5306
7.525 | -89.5322
7.5253 | -89.5333
7.5256 | -89.5339
7.5259 | -89.5343
7.5262 | -89.5365
7.5265 | -89.539
7.5268 | -89.5399
7.5271 | -89.5415
7.5274 | -89.5433
7.5278 | -89.5444
7.5281 | -89.545
7.5284 | -89.546
7.5287 | -89.5472
7.529 | -89.548
7.5293 | -89.5488
7.5296 | -89.5499
7.5299 | -89.5513
7.5302 | -89.5526
7.5305 | -89.5543
7.5308 | -89.5559
7.5312 | -89.5566
7.5315 | -89.5574
7.5318 | -89.5584
7.5321 | -89.5591
7.5324 | -89.5601
7.5327 | -89.5611
7.533 | -89.5621
7.5333 | -89.5626
7.5336 | -89.5629
7.5339 | -89.5639
7.5342 | -89.5652
7.5346 | -89.5663
7.5349 | -89.5668
7.5352 | -89.567
7.5355 | -89.568
7.5358 | -89.5691
7.5361 | -89.5711
7.5364 | -89.5733
7.5367 | -89.5748
7.537 | -89.5765
7.5373 | -89.5779
7.5376 | -89.5789
7.5379 | -89.5797
7.5383 | -89.5806
7.5386 | -89.5823
7.5389 | -89.5837
7.5392 | -89.585
7.5395 | -89.5863
7.5398 | -89.5874
7.5401 | -89.5885
7.5404 | -89.5896
7.5407 | -89.5908
7.541 | -89.592
7.5413 | -89.5929
7.5417 | -89.5939
7.542 | -89.5952
7.5423 | -89.5961
7.5426 | -89.5973
7.5429 | -89.5984
7.5432 | -89.5995
7.5435 | -89.6003
7.5438 | -89.6011
7.5441 | -89.6019
7.5444 | -89.6027
7.5447 | -89.604
7.545 | -89.6049
7.5454 | -89.6068
7.5457 | -89.6078
7.546 | -89.6094
7.5463 | -89.6105
7.5466 | -89.6116
7.5469 | -89.6128
7.5472 | -89.614
7.5475 | -89.6149
7.5478 | -89.616
7.5481 | -89.6172
7.5484 | -89.6185
7.5488 | -89.6197
7.5491 | -89.6207
7.5494 | -89.6219
7.5497 | -89.6235
7.55 | -89.6248
7.5503 | -89.6261
7.5506 | -89.6281
7.5509 | -89.6291
7.5512 | -89.6302
7.5515 | -89.6313
7.5518 | -89.6323
7.5522 | -89.6338
7.5525 | -89.6355
7.5528 | -89.6364
7.5531 | -89.6388
7.5534 | -89.6413
7.5537 | -89.6423
7.554 | -89.6441
7.5543 | -89.6448
7.5546 | -89.646
7.5549 | -89.6471
7.5552 | -89.6483
7.5555 | -89.6488
7.5559 | -89.6486
7.5562 | -89.6493
7.5565 | -89.6504
7.5568 | -89.6515
7.5571 | -89.6528
7.5574 | -89.654
7.5577 | -89.6558
7.558 | -89.6571
7.5583 | -89.6578
7.5586 | -89.6591
7.5589 | -89.6607
7.5593 | -89.662
7.5596 | -89.6632
7.5599 | -89.6637
7.5602 | -89.6643
7.5605 | -89.6649
7.5608 | -89.6661
7.5611 | -89.6674
7.5614 | -89.6686
7.5617 | -89.6696
7.562 | -89.6709
7.5623 | -89.6731
7.5626 | -89.6754
7.563 | -89.6768
7.5633 | -89.6779
7.5636 | -89.6789
7.5639 | -89.6795
7.5642 | -89.6805
7.5645 | -89.6812
7.5648 | -89.6832
7.5651 | -89.685
7.5654 | -89.6864
7.5657 | -89.6879
7.566 | -89.6889
7.5664 | -89.6899
7.5667 | -89.6909
7.567 | -89.6924
7.5673 | -89.6941
7.5676 | -89.6954
7.5679 | -89.6963
7.5682 | -89.6965
7.5685 | -89.697
7.5688 | -89.6979
7.5691 | -89.6991
7.5694 | -89.7
7.5698 | -89.701
7.5701 | -89.7019
7.5704 | -89.703
7.5707 | -89.7039
7.571 | -89.7045
7.5713 | -89.7057
7.5716 | -89.7067
7.5719 | -89.7078
7.5722 | -89.7091
7.5725 | -89.7099
7.5728 | -89.711
7.5731 | -89.7123
7.5735 | -89.7138
7.5738 | -89.7154
7.5741 | -89.7173
7.5744 | -89.7189
7.5747 | -89.7205
7.575 | -89.7213
7.5753 | -89.7216
7.5756 | -89.7228
7.5759 | -89.7239
7.5762 | -89.7251
7.5765 | -89.7262
7.5769 | -89.7271
7.5772 | -89.7281
7.5775 | -89.7291
7.5778 | -89.73
7.5781 | -89.731
7.5784 | -89.732
7.5787 | -89.7331
7.579 | -89.7343
7.5793 | -89.7355
7.5796 | -89.7365
7.5799 | -89.7374
7.5802 | -89.738
7.5806 | -89.7393
7.5809 | -89.741
7.5812 | -89.7421
7.5815 | -89.7445
7.5818 | -89.7459
7.5821 | -89.7462
7.5824 | -89.7467
7.5827 | -89.7477
7.583 | -89.7488
7.5833 | -89.7501
7.5836 | -89.7511
7.584 | -89.7524
7.5843 | -89.7543
7.5846 | -89.7556
7.5849 | -89.7569
7.5852 | -89.7581
7.5855 | -89.7596
7.5858 | -89.7611
7.5861 | -89.7624
7.5864 | -89.7637
7.5867 | -89.7649
7.587 | -89.7659
7.5873 | -89.7669
7.5877 | -89.7672
7.588 | -89.7681
7.5883 | -89.7692
7.5886 | -89.7701
7.5889 | -89.7711
7.5892 | -89.7721
7.5895 | -89.7729
7.5898 | -89.7734
7.5901 | -89.7735
7.5904 | -89.7739
7.5907 | -89.7749
7.5911 | -89.7764
7.5914 | -89.7783
7.5917 | -89.7796
7.592 | -89.7808
7.5923 | -89.7818
7.5926 | -89.7829
7.5929 | -89.7844
7.5932 | -89.7865
7.5935 | -89.7881
7.5938 | -89.7888
7.5941 | -89.7902
7.5945 | -89.7913
7.5948 | -89.7923
7.5951 | -89.7936
7.5954 | -89.7947
7.5957 | -89.7958
7.596 | -89.7968
7.5963 | -89.798
7.5966 | -89.799
7.5969 | -89.7996
7.5972 | -89.8007
7.5975 | -89.8017
7.5978 | -89.8023
7.5982 | -89.8033
7.5985 | -89.8044
7.5988 | -89.805
7.5991 | -89.8054
7.5994 | -89.8062
7.5997 | -89.8072
7.6 | -89.8082
7.6003 | -89.8092
7.6006 | -89.8105
7.6009 | -89.8117
7.6012 | -89.8125
7.6016 | -89.8129
7.6019 | -89.8137
7.6022 | -89.8151
7.6025 | -89.8169
7.6028 | -89.8175
7.6031 | -89.8177
7.6034 | -89.8172
7.6037 | -89.8172
7.604 | -89.8185
7.6043 | -89.8195
7.6046 | -89.8219
7.6049 | -89.8231
7.6053 | -89.8245
7.6056 | -89.8279
7.6059 | -89.8306
7.6062 | -89.8319
7.6065 | -89.8327
7.6068 | -89.8333
7.6071 | -89.8348
7.6074 | -89.8358
7.6077 | -89.8365
7.608 | -89.8374
7.6083 | -89.8382
7.6087 | -89.839
7.609 | -89.8403
7.6093 | -89.8412
7.6096 | -89.842
7.6099 | -89.8431
7.6102 | -89.8436
7.6105 | -89.8447
7.6108 | -89.8458
7.6111 | -89.847
7.6114 | -89.8478
7.6117 | -89.8487
7.6121 | -89.8503
7.6124 | -89.8514
7.6127 | -89.8523
7.613 | -89.8532
7.6133 | -89.8542
7.6136 | -89.855
7.6139 | -89.8558
7.6142 | -89.8568
7.6145 | -89.8578
7.6148 | -89.8583
7.6151 | -89.8592
7.6154 | -89.8611
7.6158 | -89.8624
7.6161 | -89.8629
7.6164 | -89.8637
7.6167 | -89.865
7.617 | -89.8662
7.6173 | -89.8676
7.6176 | -89.8688
7.6179 | -89.8694
7.6182 | -89.8703
7.6185 | -89.8713
7.6188 | -89.8718
7.6192 | -89.8723
7.6195 | -89.8731
7.6198 | -89.8739
7.6201 | -89.8753
7.6204 | -89.877
7.6207 | -89.8782
7.621 | -89.8794
7.6213 | -89.88
7.6216 | -89.8807
7.6219 | -89.8817
7.6222 | -89.8825
7.6226 | -89.8825
7.6229 | -89.8827
7.6232 | -89.8838
7.6235 | -89.8847
7.6238 | -89.8859
7.6241 | -89.8869
7.6244 | -89.8877
7.6247 | -89.8887
7.625 | -89.8896
7.6253 | -89.8898
7.6256 | -89.8904
7.6259 | -89.8915
7.6263 | -89.8926
7.6266 | -89.8935
7.6269 | -89.8943
7.6272 | -89.8955
7.6275 | -89.8966
7.6278 | -89.8974
7.6281 | -89.8983
7.6284 | -89.8996
7.6287 | -89.9004
7.629 | -89.9016
7.6293 | -89.9029
7.6297 | -89.904
7.63 | -89.9046
7.6303 | -89.9047
7.6306 | -89.9044
7.6309 | -89.905
7.6312 | -89.9059
7.6315 | -89.9067
7.6318 | -89.9076
7.6321 | -89.9085
7.6324 | -89.909
7.6327 | -89.9095
7.633 | -89.9103
7.6334 | -89.9116
7.6337 | -89.9131
7.634 | -89.9141
7.6343 | -89.9153
7.6346 | -89.9161
7.6349 | -89.9167
7.6352 | -89.9179
7.6355 | -89.9189
7.6358 | -89.9211
7.6361 | -89.9229
7.6364 | -89.9239
7.6368 | -89.9247
7.6371 | -89.9251
7.6374 | -89.9256
7.6377 | -89.9272
7.638 | -89.9279
7.6383 | -89.9298
7.6386 | -89.932
7.6389 | -89.9325
7.6392 | -89.9334
7.6395 | -89.9344
7.6398 | -89.935
7.6402 | -89.9354
7.6405 | -89.9357
7.6408 | -89.9365
7.6411 | -89.9388
7.6414 | -89.9409
7.6417 | -89.9419
7.642 | -89.9423
7.6423 | -89.9428
7.6426 | -89.9436
7.6429 | -89.9451
7.6432 | -89.946
7.6435 | -89.9471
7.6439 | -89.9481
7.6442 | -89.9493
7.6445 | -89.9499
7.6448 | -89.9506
7.6451 | -89.951
7.6454 | -89.9512
7.6457 | -89.9522
7.646 | -89.9535
7.6463 | -89.9544
7.6466 | -89.9549
7.6469 | -89.9553
7.6473 | -89.9561
7.6476 | -89.9567
7.6479 | -89.9577
7.6482 | -89.9585
7.6485 | -89.9585
7.6488 | -89.9586
7.6491 | -89.9594
7.6494 | -89.9599
7.6497 | -89.9605
7.65 | -89.9614
7.6503 | -89.9618
7.6506 | -89.9613
7.651 | -89.9622
7.6513 | -89.9634
7.6516 | -89.9644
7.6519 | -89.9653
7.6522 | -89.966
7.6525 | -89.9659
7.6528 | -89.9666
7.6531 | -89.9674
7.6534 | -89.9678
7.6537 | -89.9684
7.654 | -89.97
7.6544 | -89.973
7.6547 | -89.9744
7.655 | -89.9748
7.6553 | -89.9759
7.6556 | -89.9774
7.6559 | -89.9787
7.6562 | -89.9794
7.6565 | -89.9803
7.6568 | -89.9813
7.6571 | -89.982
7.6574 | -89.9823
7.6577 | -89.9823
7.6581 | -89.9825
7.6584 | -89.9832
7.6587 | -89.9844
7.659 | -89.9852
7.6593 | -89.9861
7.6596 | -89.9866
7.6599 | -89.9874
7.6602 | -89.9873
7.6605 | -89.9886
7.6608 | -89.9901
7.6611 | -89.9913
7.6615 | -89.9918
7.6618 | -89.9924
7.6621 | -89.9928
7.6624 | -89.9936
7.6627 | -89.9945
7.663 | -89.9951
7.6633 | -89.9958
7.6636 | -89.9964
7.6639 | -89.9966
7.6642 | -89.9971
7.6645 | -89.9979
7.6649 | -89.9986
7.6652 | -89.9992
7.6655 | -89.9996
7.6658 | -89.9999
7.6661 | -90.0011
7.6664 | -90.0032
7.6667 | -90.0048
7.667 | -90.005
7.6673 | -90.0056
7.6676 | -90.0065
7.6679 | -90.0079
7.6682 | -90.0087
7.6686 | -90.0088
7.6689 | -90.0098
7.6692 | -90.011
7.6695 | -90.0119
7.6698 | -90.0125
7.6701 | -90.0133
7.6704 | -90.0144
7.6707 | -90.0156
7.671 | -90.0162
7.6713 | -90.0167
7.6716 | -90.0173
7.672 | -90.0178
7.6723 | -90.0186
7.6726 | -90.02
7.6729 | -90.021
7.6732 | -90.0209
7.6735 | -90.0212
7.6738 | -90.022
7.6741 | -90.0228
7.6744 | -90.0238
7.6747 | -90.0244
7.675 | -90.0258
7.6753 | -90.0276
7.6757 | -90.0281
7.676 | -90.0283
7.6763 | -90.0291
7.6766 | -90.0296
7.6769 | -90.0304
7.6772 | -90.0319
7.6775 | -90.0331
7.6778 | -90.0335
7.6781 | -90.0337
7.6784 | -90.0346
7.6787 | -90.0351
7.6791 | -90.035
7.6794 | -90.0357
7.6797 | -90.0368
7.68 | -90.0371
7.6803 | -90.0369
7.6806 | -90.0371
7.6809 | -90.0379
7.6812 | -90.0386
7.6815 | -90.0391
7.6818 | -90.0392
7.6821 | -90.0384
7.6825 | -90.0388
7.6828 | -90.0396
7.6831 | -90.0402
7.6834 | -90.0407
7.6837 | -90.0415
7.684 | -90.042
7.6843 | -90.0428
7.6846 | -90.0435
7.6849 | -90.044
7.6852 | -90.0444
7.6855 | -90.0455
7.6858 | -90.0468
7.6862 | -90.0476
7.6865 | -90.0482
7.6868 | -90.0491
7.6871 | -90.0498
7.6874 | -90.0506
7.6877 | -90.0516
7.688 | -90.0525
7.6883 | -90.0539
7.6886 | -90.0545
7.6889 | -90.0546
7.6892 | -90.0547
7.6896 | -90.0555
7.6899 | -90.0563
7.6902 | -90.0569
7.6905 | -90.0573
7.6908 | -90.0573
7.6911 | -90.0572
7.6914 | -90.0574
7.6917 | -90.0579
7.692 | -90.0585
7.6923 | -90.059
7.6926 | -90.0598
7.6929 | -90.0603
7.6933 | -90.061
7.6936 | -90.0613
7.6939 | -90.0619
7.6942 | -90.063
7.6945 | -90.0636
7.6948 | -90.0643
7.6951 | -90.0647
7.6954 | -90.0656
7.6957 | -90.067
7.696 | -90.0684
7.6963 | -90.0693
7.6967 | -90.0699
7.697 | -90.0697
7.6973 | -90.0705
7.6976 | -90.0715
7.6979 | -90.0721
7.6982 | -90.0723
7.6985 | -90.0727
7.6988 | -90.073
7.6991 | -90.0738
7.6994 | -90.0744
7.6997 | -90.0758
7.7001 | -90.0765
7.7004 | -90.0765
7.7007 | -90.0772
7.701 | -90.0788
7.7013 | -90.0801
7.7016 | -90.0806
7.7019 | -90.0811
7.7022 | -90.0821
7.7025 | -90.0826
7.7028 | -90.0829
7.7031 | -90.0833
7.7034 | -90.0839
7.7038 | -90.0849
7.7041 | -90.0856
7.7044 | -90.086
7.7047 | -90.0865
7.705 | -90.0869
7.7053 | -90.0874
7.7056 | -90.0881
7.7059 | -90.089
7.7062 | -90.0898
7.7065 | -90.0909
7.7068 | -90.0914
7.7072 | -90.0927
7.7075 | -90.0943
7.7078 | -90.095
7.7081 | -90.0953
7.7084 | -90.0958
7.7087 | -90.0968
7.709 | -90.0981
7.7093 | -90.0988
7.7096 | -90.0997
7.7099 | -90.1007
7.7102 | -90.102
7.7105 | -90.1029
7.7109 | -90.1035
7.7112 | -90.1034
7.7115 | -90.1036
7.7118 | -90.1039
7.7121 | -90.1048
7.7124 | -90.1063
7.7127 | -90.1074
7.713 | -90.1083
7.7133 | -90.1085
7.7136 | -90.1094
7.7139 | -90.1096
7.7143 | -90.1101
7.7146 | -90.1106
7.7149 | -90.111
7.7152 | -90.1116
7.7155 | -90.1124
7.7158 | -90.1128
7.7161 | -90.1132
7.7164 | -90.1138
7.7167 | -90.1143
7.717 | -90.1148
7.7173 | -90.1152
7.7176 | -90.1157
7.718 | -90.1158
7.7183 | -90.1161
7.7186 | -90.1164
7.7189 | -90.1165
7.7192 | -90.1165
7.7195 | -90.1167
7.7198 | -90.1176
7.7201 | -90.1182
7.7204 | -90.119
7.7207 | -90.119
7.721 | -90.1191
7.7214 | -90.1199
7.7217 | -90.1204
7.722 | -90.121
7.7223 | -90.1216
7.7226 | -90.122
7.7229 | -90.1225
7.7232 | -90.1231
7.7235 | -90.1241
7.7238 | -90.1241
7.7241 | -90.125
7.7244 | -90.1257
7.7248 | -90.1267
7.7251 | -90.1273
7.7254 | -90.1275
7.7257 | -90.128
7.726 | -90.1285
7.7263 | -90.1284
7.7266 | -90.129
7.7269 | -90.1293
7.7272 | -90.1292
7.7275 | -90.1283
7.7278 | -90.1276
7.7281 | -90.1277
7.7285 | -90.1283
7.7288 | -90.129
7.7291 | -90.1299
7.7294 | -90.1303
7.7297 | -90.1305
7.73 | -90.1314
7.7303 | -90.1316
7.7306 | -90.132
7.7309 | -90.1328
7.7312 | -90.1355
7.7315 | -90.137
7.7319 | -90.1377
7.7322 | -90.1387
7.7325 | -90.1393
7.7328 | -90.1394
7.7331 | -90.14
7.7334 | -90.1409
7.7337 | -90.1416
7.734 | -90.1432
7.7343 | -90.1443
7.7346 | -90.1449
7.7349 | -90.1454
7.7352 | -90.1465
7.7356 | -90.1468
7.7359 | -90.1472
7.7362 | -90.1478
7.7365 | -90.1485
7.7368 | -90.1488
7.7371 | -90.1495
7.7374 | -90.15
7.7377 | -90.1508
7.738 | -90.152
7.7383 | -90.153
7.7386 | -90.1532
7.739 | -90.1537
7.7393 | -90.1542
7.7396 | -90.1547
7.7399 | -90.1551
7.7402 | -90.1558
7.7405 | -90.1563
7.7408 | -90.1567
7.7411 | -90.1573
7.7414 | -90.1566
7.7417 | -90.157
7.742 | -90.1573
7.7424 | -90.158
7.7427 | -90.1587
7.743 | -90.1591
7.7433 | -90.1598
7.7436 | -90.1602
7.7439 | -90.1613
7.7442 | -90.1624
7.7445 | -90.163
7.7448 | -90.1634
7.7451 | -90.1637
7.7454 | -90.1637
7.7457 | -90.1641
7.7461 | -90.1635
7.7464 | -90.1648
7.7467 | -90.1655
7.747 | -90.1654
7.7473 | -90.1657
7.7476 | -90.1661
7.7479 | -90.1668
7.7482 | -90.1676
7.7485 | -90.1684
7.7488 | -90.1688
7.7491 | -90.1693
7.7495 | -90.1697
7.7498 | -90.1702
7.7501 | -90.1713
7.7504 | -90.1711
7.7507 | -90.172
7.751 | -90.1721
7.7513 | -90.1725
7.7516 | -90.1736
7.7519 | -90.1744
7.7522 | -90.1744
7.7525 | -90.1747
7.7528 | -90.1755
7.7532 | -90.1765
7.7535 | -90.1772
7.7538 | -90.1779
7.7541 | -90.1777
7.7544 | -90.1782
7.7547 | -90.1786
7.755 | -90.179
7.7553 | -90.1806
7.7556 | -90.1817
7.7559 | -90.1822
7.7562 | -90.1826
7.7566 | -90.1829
7.7569 | -90.1832
7.7572 | -90.1836
7.7575 | -90.1836
7.7578 | -90.1843
7.7581 | -90.1857
7.7584 | -90.1863
7.7587 | -90.1856
7.759 | -90.185
7.7593 | -90.1852
7.7596 | -90.185
7.76 | -90.1859
7.7603 | -90.1873
7.7606 | -90.1881
7.7609 | -90.1884
7.7612 | -90.1885
7.7615 | -90.1884
7.7618 | -90.188
7.7621 | -90.1889
7.7624 | -90.1896
7.7627 | -90.1904
7.763 | -90.1907
7.7633 | -90.191
7.7637 | -90.1919
7.764 | -90.1924
7.7643 | -90.1927
7.7646 | -90.193
7.7649 | -90.1932
7.7652 | -90.1941
7.7655 | -90.1946
7.7658 | -90.1953
7.7661 | -90.1962
7.7664 | -90.1967
7.7667 | -90.197
7.7671 | -90.1974
7.7674 | -90.198
7.7677 | -90.1985
7.768 | -90.1989
7.7683 | -90.1996
7.7686 | -90.1998
7.7689 | -90.1994
7.7692 | -90.1993
7.7695 | -90.1988
7.7698 | -90.1985
7.7701 | -90.1986
7.7705 | -90.1985
7.7708 | -90.199
7.7711 | -90.1997
7.7714 | -90.2007
7.7717 | -90.2015
7.772 | -90.2019
7.7723 | -90.2024
7.7726 | -90.2033
7.7729 | -90.2039
7.7732 | -90.2043
7.7735 | -90.2048
7.7738 | -90.2057
7.7742 | -90.2062
7.7745 | -90.2064
7.7748 | -90.2076
7.7751 | -90.2086
7.7754 | -90.2088
7.7757 | -90.2091
7.776 | -90.2104
7.7763 | -90.2116
7.7766 | -90.2121
7.7769 | -90.213
7.7772 | -90.2137
7.7776 | -90.2141
7.7779 | -90.2144
7.7782 | -90.2153
7.7785 | -90.216
7.7788 | -90.2165
7.7791 | -90.2173
7.7794 | -90.2177
7.7797 | -90.2181
7.78 | -90.2187
7.7803 | -90.2198
7.7806 | -90.2206
7.7809 | -90.2216
7.7813 | -90.2217
7.7816 | -90.2216
7.7819 | -90.2219
7.7822 | -90.2225
7.7825 | -90.2231
7.7828 | -90.2237
7.7831 | -90.2247
7.7834 | -90.2258
7.7837 | -90.226
7.784 | -90.2261
7.7843 | -90.2267
7.7847 | -90.2266
7.785 | -90.2269
7.7853 | -90.227
7.7856 | -90.2274
7.7859 | -90.2278
7.7862 | -90.228
7.7865 | -90.2285
7.7868 | -90.2286
7.7871 | -90.2293
7.7874 | -90.2302
7.7877 | -90.2306
7.788 | -90.2312
7.7884 | -90.2324
7.7887 | -90.2327
7.789 | -90.2332
7.7893 | -90.2333
7.7896 | -90.234
7.7899 | -90.2348
7.7902 | -90.2354
7.7905 | -90.2358
7.7908 | -90.2362
7.7911 | -90.2364
7.7914 | -90.2369
7.7918 | -90.2373
7.7921 | -90.2377
7.7924 | -90.2382
7.7927 | -90.2388
7.793 | -90.2395
7.7933 | -90.2401
7.7936 | -90.2409
7.7939 | -90.241
7.7942 | -90.2414
7.7945 | -90.2418
7.7948 | -90.2417
7.7952 | -90.2418
7.7955 | -90.2423
7.7958 | -90.2426
7.7961 | -90.2431
7.7964 | -90.2435
7.7967 | -90.2437
7.797 | -90.244
7.7973 | -90.2444
7.7976 | -90.2447
7.7979 | -90.245
7.7982 | -90.2454
7.7985 | -90.2457
7.7989 | -90.2459
7.7992 | -90.2464
7.7995 | -90.2469
7.7998 | -90.2472
7.8001 | -90.2475
7.8004 | -90.248
7.8007 | -90.2488
7.801 | -90.2495
7.8013 | -90.2498
7.8016 | -90.2499
7.8019 | -90.2503
7.8023 | -90.25
7.8026 | -90.2501
7.8029 | -90.2504
7.8032 | -90.2505
7.8035 | -90.2511
7.8038 | -90.251
7.8041 | -90.2516
7.8044 | -90.2517
7.8047 | -90.2519
7.805 | -90.2527
7.8053 | -90.2533
7.8056 | -90.2535
7.806 | -90.2539
7.8063 | -90.2543
7.8066 | -90.2546
7.8069 | -90.255
7.8072 | -90.2557
7.8075 | -90.2561
7.8078 | -90.2559
7.8081 | -90.2561
7.8084 | -90.2565
7.8087 | -90.257
7.809 | -90.2565
7.8094 | -90.2573
7.8097 | -90.258
7.81 | -90.2584
7.8103 | -90.2592
7.8106 | -90.2607
7.8109 | -90.2625
7.8112 | -90.2631
7.8115 | -90.2632
7.8118 | -90.2629
7.8121 | -90.2628
7.8124 | -90.2633
7.8128 | -90.2639
7.8131 | -90.2638
7.8134 | -90.264
7.8137 | -90.264
7.814 | -90.2643
7.8143 | -90.265
7.8146 | -90.2657
7.8149 | -90.2662
7.8152 | -90.2667
7.8155 | -90.2672
7.8158 | -90.2676
7.8161 | -90.2677
7.8165 | -90.2678
7.8168 | -90.2684
7.8171 | -90.2689
7.8174 | -90.2682
7.8177 | -90.2671
7.818 | -90.2666
7.8183 | -90.2667
7.8186 | -90.267
7.8189 | -90.2674
7.8192 | -90.2677
7.8195 | -90.268
7.8199 | -90.2684
7.8202 | -90.2688
7.8205 | -90.2696
7.8208 | -90.2712
7.8211 | -90.272
7.8214 | -90.2718
7.8217 | -90.2722
7.822 | -90.2726
7.8223 | -90.2736
7.8226 | -90.275
7.8229 | -90.276
7.8232 | -90.2769
7.8236 | -90.2774
7.8239 | -90.2776
7.8242 | -90.2777
7.8245 | -90.2786
7.8248 | -90.2793
7.8251 | -90.28
7.8254 | -90.2804
7.8257 | -90.2803
7.826 | -90.2805
7.8263 | -90.2813
7.8266 | -90.283
7.827 | -90.2832
7.8273 | -90.2844
7.8276 | -90.2851
7.8279 | -90.2855
7.8282 | -90.2858
7.8285 | -90.2864
7.8288 | -90.2867
7.8291 | -90.2871
7.8294 | -90.2876
7.8297 | -90.2881
7.83 | -90.2884
7.8304 | -90.2888
7.8307 | -90.2894
7.831 | -90.2896
7.8313 | -90.289
7.8316 | -90.2882
7.8319 | -90.2882
7.8322 | -90.2886
7.8325 | -90.289
7.8328 | -90.2889
7.8331 | -90.2895
7.8334 | -90.2909
7.8337 | -90.2921
7.8341 | -90.2926
7.8344 | -90.2936
7.8347 | -90.2953
7.835 | -90.2963
7.8353 | -90.2973
7.8356 | -90.2984
7.8359 | -90.2991
7.8362 | -90.2994
7.8365 | -90.2997
7.8368 | -90.3007
7.8371 | -90.3014
7.8375 | -90.3017
7.8378 | -90.3018
7.8381 | -90.3021
7.8384 | -90.3022
7.8387 | -90.3025
7.839 | -90.3029
7.8393 | -90.3036
7.8396 | -90.3035
7.8399 | -90.3037
7.8402 | -90.3043
7.8405 | -90.3048
7.8408 | -90.3052
7.8412 | -90.3055
7.8415 | -90.3057
7.8418 | -90.3064
7.8421 | -90.3072
7.8424 | -90.3077
7.8427 | -90.308
7.843 | -90.3085
7.8433 | -90.3091
7.8436 | -90.3094
7.8439 | -90.3101
7.8442 | -90.3106
7.8446 | -90.3106
7.8449 | -90.3107
7.8452 | -90.3109
7.8455 | -90.3111
7.8458 | -90.3116
7.8461 | -90.3124
7.8464 | -90.3127
7.8467 | -90.3133
7.847 | -90.3143
7.8473 | -90.3149
7.8476 | -90.3163
7.8479 | -90.3173
7.8483 | -90.3182
7.8486 | -90.3182
7.8489 | -90.3183
7.8492 | -90.3184
7.8495 | -90.3188
7.8498 | -90.319
7.8501 | -90.3194
7.8504 | -90.3199
7.8507 | -90.3199
7.851 | -90.3202
7.8513 | -90.3205
7.8517 | -90.3211
7.852 | -90.3215
7.8523 | -90.3219
7.8526 | -90.3224
7.8529 | -90.3228
7.8532 | -90.3235
7.8535 | -90.3242
7.8538 | -90.3245
7.8541 | -90.3241
7.8544 | -90.3241
7.8547 | -90.3246
7.8551 | -90.3248
7.8554 | -90.3252
7.8557 | -90.3253
7.856 | -90.3255
7.8563 | -90.3258
7.8566 | -90.3263
7.8569 | -90.3267
7.8572 | -90.3272
7.8575 | -90.3274
7.8578 | -90.328
7.8581 | -90.3292
7.8584 | -90.3298
7.8588 | -90.3303
7.8591 | -90.3308
7.8594 | -90.3311
7.8597 | -90.3317
7.86 | -90.3322
7.8603 | -90.3325
7.8606 | -90.3329
7.8609 | -90.3333
7.8612 | -90.334
7.8615 | -90.3344
7.8618 | -90.3349
7.8622 | -90.3354
7.8625 | -90.3357
7.8628 | -90.3357
7.8631 | -90.3359
7.8634 | -90.3358
7.8637 | -90.3358
7.864 | -90.3362
7.8643 | -90.3366
7.8646 | -90.3372
7.8649 | -90.3377
7.8652 | -90.3385
7.8655 | -90.3386
7.8659 | -90.3389
7.8662 | -90.3392
7.8665 | -90.3401
7.8668 | -90.3414
7.8671 | -90.3422
7.8674 | -90.3426
7.8677 | -90.343
7.868 | -90.3433
7.8683 | -90.3437
7.8686 | -90.3441
7.8689 | -90.3443
7.8693 | -90.3448
7.8696 | -90.3452
7.8699 | -90.3455
7.8702 | -90.3462
7.8705 | -90.3464
7.8708 | -90.3469
7.8711 | -90.3476
7.8714 | -90.3483
7.8717 | -90.3485
7.872 | -90.3485
7.8723 | -90.3493
7.8727 | -90.35
7.873 | -90.3523
7.8733 | -90.3534
7.8736 | -90.3538
7.8739 | -90.3547
7.8742 | -90.3556
7.8745 | -90.3569
7.8748 | -90.3579
7.8751 | -90.3587
7.8754 | -90.3591
7.8757 | -90.36
7.876 | -90.3601
7.8764 | -90.3601
7.8767 | -90.3603
7.877 | -90.3609
7.8773 | -90.3618
7.8776 | -90.3622
7.8779 | -90.3628
7.8782 | -90.3635
7.8785 | -90.3645
7.8788 | -90.3653
7.8791 | -90.3655
7.8794 | -90.3661
7.8798 | -90.3665
7.8801 | -90.3671
7.8804 | -90.3675
7.8807 | -90.3673
7.881 | -90.3676
7.8813 | -90.3681
7.8816 | -90.3684
7.8819 | -90.3688
7.8822 | -90.3698
7.8825 | -90.3702
7.8828 | -90.3699
7.8831 | -90.3701
7.8835 | -90.3704
7.8838 | -90.3706
7.8841 | -90.3715
7.8844 | -90.372
7.8847 | -90.3725
7.885 | -90.3727
7.8853 | -90.3734
7.8856 | -90.374
7.8859 | -90.3748
7.8862 | -90.3754
7.8865 | -90.3761
7.8869 | -90.3767
7.8872 | -90.377
7.8875 | -90.3774
7.8878 | -90.3779
7.8881 | -90.3783
7.8884 | -90.3786
7.8887 | -90.3787
7.889 | -90.3787
7.8893 | -90.3789
7.8896 | -90.3797
7.8899 | -90.3809
7.8903 | -90.3822
7.8906 | -90.3827
7.8909 | -90.3832
7.8912 | -90.3837
7.8915 | -90.3842
7.8918 | -90.3851
7.8921 | -90.3856
7.8924 | -90.3861
7.8927 | -90.3861
7.893 | -90.3866
7.8933 | -90.3881
7.8936 | -90.3894
7.894 | -90.3899
7.8943 | -90.3903
7.8946 | -90.3906
7.8949 | -90.3908
7.8952 | -90.3907
7.8955 | -90.3902
7.8958 | -90.3904
7.8961 | -90.3909
7.8964 | -90.3917
7.8967 | -90.3923
7.897 | -90.3928
7.8974 | -90.3934
7.8977 | -90.3938
7.898 | -90.3943
7.8983 | -90.3953
7.8986 | -90.3958
7.8989 | -90.3971
7.8992 | -90.3983
7.8995 | -90.3988
7.8998 | -90.3995
7.9001 | -90.3998
7.9004 | -90.4006
7.9008 | -90.4011
7.9011 | -90.4017
7.9014 | -90.4022
7.9017 | -90.4035
7.902 | -90.4045
7.9023 | -90.406
7.9026 | -90.4067
7.9029 | -90.4075
7.9032 | -90.4081
7.9035 | -90.4082
7.9038 | -90.4086
7.9041 | -90.4097
7.9045 | -90.4103
7.9048 | -90.4109
7.9051 | -90.4112
7.9054 | -90.4124
7.9057 | -90.4135
7.906 | -90.4133
7.9063 | -90.4128
7.9066 | -90.413
7.9069 | -90.4136
7.9072 | -90.4139
7.9075 | -90.4141
7.9079 | -90.4144
7.9082 | -90.415
7.9085 | -90.4158
7.9088 | -90.4167
7.9091 | -90.4172
7.9094 | -90.4175
7.9097 | -90.4176
7.91 | -90.418
7.9103 | -90.4186
7.9106 | -90.4194
7.9109 | -90.4199
7.9112 | -90.4203
7.9116 | -90.4205
7.9119 | -90.4211
7.9122 | -90.4221
7.9125 | -90.423
7.9128 | -90.4235
7.9131 | -90.4239
7.9134 | -90.4242
7.9137 | -90.4247
7.914 | -90.4253
7.9143 | -90.4254
7.9146 | -90.4258
7.915 | -90.4264
7.9153 | -90.427
7.9156 | -90.4275
7.9159 | -90.428
7.9162 | -90.4277
7.9165 | -90.4289
7.9168 | -90.4292
7.9171 | -90.4297
7.9174 | -90.4309
7.9177 | -90.4316
7.918 | -90.432
7.9183 | -90.4327
7.9187 | -90.4333
7.919 | -90.4332
7.9193 | -90.4329
7.9196 | -90.4336
7.9199 | -90.4345
7.9202 | -90.4355
7.9205 | -90.4358
7.9208 | -90.4361
7.9211 | -90.4367
7.9214 | -90.4373
7.9217 | -90.4379
7.9221 | -90.4384
7.9224 | -90.4391
7.9227 | -90.4399
7.923 | -90.4408
7.9233 | -90.4425
7.9236 | -90.4431
7.9239 | -90.4438
7.9242 | -90.4444
7.9245 | -90.445
7.9248 | -90.4468
7.9251 | -90.4478
7.9255 | -90.4481
7.9258 | -90.449
7.9261 | -90.4495
7.9264 | -90.4498
7.9267 | -90.4507
7.927 | -90.4513
7.9273 | -90.451
7.9276 | -90.4508
7.9279 | -90.4519
7.9282 | -90.4528
7.9285 | -90.4536
7.9288 | -90.4539
7.9292 | -90.4542
7.9295 | -90.4544
7.9298 | -90.4542
7.9301 | -90.4544
7.9304 | -90.455
7.9307 | -90.4556
7.931 | -90.4563
7.9313 | -90.457
7.9316 | -90.4578
7.9319 | -90.4584
7.9322 | -90.4589
7.9326 | -90.4593
7.9329 | -90.4599
7.9332 | -90.4608
7.9335 | -90.4615
7.9338 | -90.4624
7.9341 | -90.4636
7.9344 | -90.4649
7.9347 | -90.4658
7.935 | -90.4668
7.9353 | -90.4686
7.9356 | -90.4703
7.9359 | -90.4708
7.9363 | -90.4715
7.9366 | -90.472
7.9369 | -90.4722
7.9372 | -90.472
7.9375 | -90.4716
7.9378 | -90.472
7.9381 | -90.4727
7.9384 | -90.4734
7.9387 | -90.474
7.939 | -90.4749
7.9393 | -90.4758
7.9397 | -90.4765
7.94 | -90.4776
7.9403 | -90.4792
7.9406 | -90.4809
7.9409 | -90.4816
7.9412 | -90.4822
7.9415 | -90.4827
7.9418 | -90.4831
7.9421 | -90.4836
7.9424 | -90.4842
7.9427 | -90.4851
7.9431 | -90.4862
7.9434 | -90.4865
7.9437 | -90.4865
7.944 | -90.4874
7.9443 | -90.4881
7.9446 | -90.4885
7.9449 | -90.4889
7.9452 | -90.4895
7.9455 | -90.4902
7.9458 | -90.4904
7.9461 | -90.4908
7.9464 | -90.4918
7.9468 | -90.4922
7.9471 | -90.4928
7.9474 | -90.4935
7.9477 | -90.4946
7.948 | -90.4953
7.9483 | -90.4958
7.9486 | -90.4964
7.9489 | -90.4973
7.9492 | -90.4987
7.9495 | -90.4985
7.9498 | -90.4986
7.9502 | -90.4993
7.9505 | -90.4997
7.9508 | -90.5006
7.9511 | -90.5016
7.9514 | -90.5022
7.9517 | -90.5034
7.952 | -90.5043
7.9523 | -90.5053
7.9526 | -90.506
7.9529 | -90.5066
7.9532 | -90.5075
7.9535 | -90.5082
7.9539 | -90.5089
7.9542 | -90.5099
7.9545 | -90.5113
7.9548 | -90.512
7.9551 | -90.5126
7.9554 | -90.5133
7.9557 | -90.5139
7.956 | -90.5146
7.9563 | -90.5153
7.9566 | -90.5158
7.9569 | -90.5167
7.9573 | -90.5176
7.9576 | -90.5179
7.9579 | -90.5186
7.9582 | -90.5189
7.9585 | -90.5192
7.9588 | -90.5197
7.9591 | -90.5204
7.9594 | -90.5205
7.9597 | -90.5211
7.96 | -90.5221
7.9603 | -90.523
7.9607 | -90.5238
7.961 | -90.5244
7.9613 | -90.5253
7.9616 | -90.526
7.9619 | -90.5265
7.9622 | -90.5273
7.9625 | -90.528
7.9628 | -90.5283
7.9631 | -90.5288
7.9634 | -90.5298
7.9637 | -90.532
7.964 | -90.533
7.9644 | -90.5336
7.9647 | -90.5344
7.965 | -90.535
7.9653 | -90.5356
7.9656 | -90.5357
7.9659 | -90.5359
7.9662 | -90.5365
7.9665 | -90.5375
7.9668 | -90.5384
7.9671 | -90.5393
7.9674 | -90.54
7.9678 | -90.5408
7.9681 | -90.5417
7.9684 | -90.5429
7.9687 | -90.5438
7.969 | -90.5451
7.9693 | -90.5469
7.9696 | -90.5478
7.9699 | -90.5485
7.9702 | -90.5489
7.9705 | -90.5496
7.9708 | -90.5504
7.9711 | -90.5513
7.9715 | -90.5521
7.9718 | -90.5528
7.9721 | -90.5535
7.9724 | -90.5541
7.9727 | -90.5548
7.973 | -90.556
7.9733 | -90.5569
7.9736 | -90.5577
7.9739 | -90.5583
7.9742 | -90.5586
7.9745 | -90.5588
7.9749 | -90.5597
7.9752 | -90.5599
7.9755 | -90.5603
7.9758 | -90.561
7.9761 | -90.5617
7.9764 | -90.5622
7.9767 | -90.5627
7.977 | -90.5643
7.9773 | -90.5658
7.9776 | -90.5665
7.9779 | -90.5667
7.9783 | -90.5673
7.9786 | -90.5679
7.9789 | -90.5686
7.9792 | -90.5695
7.9795 | -90.5704
7.9798 | -90.5712
7.9801 | -90.5719
7.9804 | -90.5727
7.9807 | -90.5733
7.981 | -90.5741
7.9813 | -90.5762
7.9816 | -90.5779
7.982 | -90.5786
7.9823 | -90.5793
7.9826 | -90.5809
7.9829 | -90.5831
7.9832 | -90.5842
7.9835 | -90.585
7.9838 | -90.586
7.9841 | -90.5879
7.9844 | -90.5893
7.9847 | -90.5901
7.985 | -90.5909
7.9854 | -90.5922
7.9857 | -90.5941
7.986 | -90.5953
7.9863 | -90.5961
7.9866 | -90.597
7.9869 | -90.598
7.9872 | -90.5994
7.9875 | -90.6007
7.9878 | -90.6017
7.9881 | -90.6025
7.9884 | -90.6035
7.9887 | -90.6047
7.9891 | -90.6059
7.9894 | -90.607
7.9897 | -90.6078
7.99 | -90.6086
7.9903 | -90.6089
7.9906 | -90.609
7.9909 | -90.6101
7.9912 | -90.6113
7.9915 | -90.6126
7.9918 | -90.6134
7.9921 | -90.6145
7.9925 | -90.6162
7.9928 | -90.6182
7.9931 | -90.6192
7.9934 | -90.6197
7.9937 | -90.6204
7.994 | -90.6211
7.9943 | -90.622
7.9946 | -90.6226
7.9949 | -90.6233
7.9952 | -90.6246
7.9955 | -90.6263
7.9958 | -90.6272
7.9962 | -90.6275
7.9965 | -90.6281
7.9968 | -90.6288
7.9971 | -90.6296
7.9974 | -90.6305
7.9977 | -90.6313
7.998 | -90.6319
7.9983 | -90.6327
7.9986 | -90.6336
7.9989 | -90.6341
7.9992 | -90.6346
7.9996 | -90.6356
7.9999 | -90.6368
8.0002 | -90.6375
8.0005 | -90.6384
8.0008 | -90.6392
8.0011 | -90.6403
8.0014 | -90.6422
8.0017 | -90.6435
8.002 | -90.6442
8.0023 | -90.6446
8.0026 | -90.6453
8.003 | -90.6462
8.0033 | -90.6473
8.0036 | -90.6484
8.0039 | -90.6499
8.0042 | -90.6507
8.0045 | -90.6512
8.0048 | -90.6522
8.0051 | -90.6534
8.0054 | -90.6541
8.0057 | -90.6551
8.006 | -90.6563
8.0063 | -90.6571
8.0067 | -90.6583
8.007 | -90.6595
8.0073 | -90.6605
8.0076 | -90.6613
8.0079 | -90.6625
8.0082 | -90.6635
8.0085 | -90.6646
8.0088 | -90.6659
8.0091 | -90.667
8.0094 | -90.6683
8.0097 | -90.6693
8.0101 | -90.6701
8.0104 | -90.6709
8.0107 | -90.6724
8.011 | -90.6744
8.0113 | -90.6756
8.0116 | -90.6765
8.0119 | -90.6775
8.0122 | -90.6786
8.0125 | -90.6793
8.0128 | -90.6801
8.0131 | -90.681
8.0135 | -90.6821
8.0138 | -90.6828
8.0141 | -90.6832
8.0144 | -90.6834
8.0147 | -90.6839
8.015 | -90.6847
8.0153 | -90.6854
8.0156 | -90.6874
8.0159 | -90.6892
8.0162 | -90.6905
8.0165 | -90.6917
8.0168 | -90.6924
8.0172 | -90.6931
8.0175 | -90.6943
8.0178 | -90.6954
8.0181 | -90.6963
8.0184 | -90.6972
8.0187 | -90.6978
8.019 | -90.6988
8.0193 | -90.6997
8.0196 | -90.7004
8.0199 | -90.7011
8.0202 | -90.702
8.0206 | -90.7027
8.0209 | -90.7032
8.0212 | -90.704
8.0215 | -90.705
8.0218 | -90.7064
8.0221 | -90.7072
8.0224 | -90.7083
8.0227 | -90.7094
8.023 | -90.7104
8.0233 | -90.7113
8.0236 | -90.7124
8.0239 | -90.7133
8.0243 | -90.7143
8.0246 | -90.7151
8.0249 | -90.7152
8.0252 | -90.7165
8.0255 | -90.7186
8.0258 | -90.7192
8.0261 | -90.7197
8.0264 | -90.7212
8.0267 | -90.722
8.027 | -90.7234
8.0273 | -90.7245
8.0277 | -90.7253
8.028 | -90.7263
8.0283 | -90.7269
8.0286 | -90.7269
8.0289 | -90.7267
8.0292 | -90.7278
8.0295 | -90.7289
8.0298 | -90.7297
8.0301 | -90.7305
8.0304 | -90.7314
8.0307 | -90.7325
8.0311 | -90.7346
8.0314 | -90.736
8.0317 | -90.737
8.032 | -90.738
8.0323 | -90.7385
8.0326 | -90.7393
8.0329 | -90.7398
8.0332 | -90.7405
8.0335 | -90.7415
8.0338 | -90.7424
8.0341 | -90.7431
8.0344 | -90.7443
8.0348 | -90.7453
8.0351 | -90.7468
8.0354 | -90.748
8.0357 | -90.749
8.036 | -90.7504
8.0363 | -90.7516
8.0366 | -90.7529
8.0369 | -90.754
8.0372 | -90.7548
8.0375 | -90.756
8.0378 | -90.757
8.0382 | -90.7581
8.0385 | -90.759
8.0388 | -90.76
8.0391 | -90.7609
8.0394 | -90.762
8.0397 | -90.7628
8.04 | -90.7638
8.0403 | -90.764
8.0406 | -90.7642
8.0409 | -90.765
8.0412 | -90.7662
8.0415 | -90.7671
8.0419 | -90.7679
8.0422 | -90.768
8.0425 | -90.7683
8.0428 | -90.7685
8.0431 | -90.7688
8.0434 | -90.7693
8.0437 | -90.77
8.044 | -90.7709
8.0443 | -90.7711
8.0446 | -90.7719
8.0449 | -90.7723
8.0453 | -90.7731
8.0456 | -90.7737
8.0459 | -90.7735
8.0462 | -90.7723
8.0465 | -90.7729
8.0468 | -90.7738
8.0471 | -90.7747
8.0474 | -90.7762
8.0477 | -90.7777
8.048 | -90.7784
8.0483 | -90.7797
8.0486 | -90.7808
8.049 | -90.7816
8.0493 | -90.7823
8.0496 | -90.7828
8.0499 | -90.7835
8.0502 | -90.7846
8.0505 | -90.7855
8.0508 | -90.7861
8.0511 | -90.787
8.0514 | -90.7877
8.0517 | -90.7888
8.052 | -90.7898
8.0524 | -90.79
8.0527 | -90.7904
8.053 | -90.7916
8.0533 | -90.7916
8.0536 | -90.7923
8.0539 | -90.7937
8.0542 | -90.7951
8.0545 | -90.7959
8.0548 | -90.7968
8.0551 | -90.7976
8.0554 | -90.7988
8.0558 | -90.8
8.0561 | -90.801
8.0564 | -90.8024
8.0567 | -90.8035
8.057 | -90.8046
8.0573 | -90.8051
8.0576 | -90.8055
8.0579 | -90.807
8.0582 | -90.8089
8.0585 | -90.8098
8.0588 | -90.8108
8.0591 | -90.8117
8.0595 | -90.8125
8.0598 | -90.814
8.0601 | -90.8155
8.0604 | -90.8168
8.0607 | -90.8176
8.061 | -90.8189
8.0613 | -90.8199
8.0616 | -90.8206
8.0619 | -90.8213
8.0622 | -90.8224
8.0625 | -90.8231
8.0629 | -90.8242
8.0632 | -90.8248
8.0635 | -90.8263
8.0638 | -90.8284
8.0641 | -90.8298
8.0644 | -90.8311
8.0647 | -90.8317
8.065 | -90.8322
8.0653 | -90.8331
8.0656 | -90.8335
8.0659 | -90.8343
8.0662 | -90.835
8.0666 | -90.8365
8.0669 | -90.838
8.0672 | -90.8389
8.0675 | -90.8398
8.0678 | -90.8405
8.0681 | -90.8411
8.0684 | -90.8416
8.0687 | -90.8427
8.069 | -90.8439
8.0693 | -90.8453
8.0696 | -90.8461
8.07 | -90.847
8.0703 | -90.8477
8.0706 | -90.8483
8.0709 | -90.8497
8.0712 | -90.8508
8.0715 | -90.8518
8.0718 | -90.8521
8.0721 | -90.8526
8.0724 | -90.8532
8.0727 | -90.8539
8.073 | -90.8541
8.0734 | -90.8552
8.0737 | -90.855
8.074 | -90.8549
8.0743 | -90.8554
8.0746 | -90.8563
8.0749 | -90.8579
8.0752 | -90.8599
8.0755 | -90.861
8.0758 | -90.8619
8.0761 | -90.8629
8.0764 | -90.8635
8.0767 | -90.8646
8.0771 | -90.8653
8.0774 | -90.866
8.0777 | -90.8668
8.078 | -90.8678
8.0783 | -90.8684
8.0786 | -90.8692
8.0789 | -90.8699
8.0792 | -90.8704
8.0795 | -90.8714
8.0798 | -90.8722
8.0801 | -90.8729
8.0805 | -90.8739
8.0808 | -90.8746
8.0811 | -90.8743
8.0814 | -90.8748
8.0817 | -90.8757
8.082 | -90.8766
8.0823 | -90.8774
8.0826 | -90.878
8.0829 | -90.8781
8.0832 | -90.879
8.0835 | -90.8801
8.0838 | -90.8812
8.0842 | -90.8821
8.0845 | -90.8829
8.0848 | -90.8836
8.0851 | -90.8846
8.0854 | -90.8853
8.0857 | -90.8863
8.086 | -90.8872
8.0863 | -90.888
8.0866 | -90.8892
8.0869 | -90.8901
8.0872 | -90.8914
8.0876 | -90.8925
8.0879 | -90.8933
8.0882 | -90.8947
8.0885 | -90.8959
8.0888 | -90.8967
8.0891 | -90.8976
8.0894 | -90.8986
8.0897 | -90.8996
8.09 | -90.9007
8.0903 | -90.9017
8.0906 | -90.9025
8.091 | -90.9033
8.0913 | -90.9042
8.0916 | -90.9049
8.0919 | -90.9058
8.0922 | -90.9066
8.0925 | -90.9075
8.0928 | -90.9086
8.0931 | -90.9095
8.0934 | -90.9104
8.0937 | -90.9111
8.094 | -90.9117
8.0943 | -90.9126
8.0947 | -90.9136
8.095 | -90.9145
8.0953 | -90.9149
8.0956 | -90.9163
8.0959 | -90.9173
8.0962 | -90.918
8.0965 | -90.9193
8.0968 | -90.9196
8.0971 | -90.9202
8.0974 | -90.9215
8.0977 | -90.9223
8.0981 | -90.923
8.0984 | -90.9234
8.0987 | -90.9244
8.099 | -90.9266
8.0993 | -90.9293
8.0996 | -90.9305
8.0999 | -90.9313
8.1002 | -90.9321
8.1005 | -90.9328
8.1008 | -90.9337
8.1011 | -90.9345
8.1014 | -90.9357
8.1018 | -90.9371
8.1021 | -90.9379
8.1024 | -90.9384
8.1027 | -90.9391
8.103 | -90.94
8.1033 | -90.9415
8.1036 | -90.9429
8.1039 | -90.944
8.1042 | -90.9447
8.1045 | -90.945
8.1048 | -90.9457
8.1052 | -90.9458
8.1055 | -90.9458
8.1058 | -90.9457
8.1061 | -90.9463
8.1064 | -90.9471
8.1067 | -90.9474
8.107 | -90.9481
8.1073 | -90.9492
8.1076 | -90.95
8.1079 | -90.9514
8.1082 | -90.9518
8.1085 | -90.9524
8.1089 | -90.953
8.1092 | -90.9534
8.1095 | -90.9536
8.1098 | -90.9547
8.1101 | -90.9554
8.1104 | -90.9559
8.1107 | -90.9567
8.111 | -90.9573
8.1113 | -90.9575
8.1116 | -90.9581
8.1119 | -90.959
8.1123 | -90.9608
8.1126 | -90.9621
8.1129 | -90.9629
8.1132 | -90.9627
8.1135 | -90.9624
8.1138 | -90.9631
8.1141 | -90.9639
8.1144 | -90.9647
8.1147 | -90.9653
8.115 | -90.9665
8.1153 | -90.967
8.1157 | -90.9678
8.116 | -90.9686
8.1163 | -90.9693
8.1166 | -90.97
8.1169 | -90.971
8.1172 | -90.972
8.1175 | -90.9726
8.1178 | -90.9734
8.1181 | -90.9744
8.1184 | -90.9752
8.1187 | -90.9757
8.119 | -90.9764
8.1194 | -90.9771
8.1197 | -90.9786
8.12 | -90.9805
8.1203 | -90.9814
8.1206 | -90.9822
8.1209 | -90.9829
8.1212 | -90.9833
8.1215 | -90.9834
8.1218 | -90.9841
8.1221 | -90.9846
8.1224 | -90.9854
8.1228 | -90.9864
8.1231 | -90.9872
8.1234 | -90.988
8.1237 | -90.9888
8.124 | -90.9893
8.1243 | -90.9899
8.1246 | -90.9905
8.1249 | -90.9916
8.1252 | -90.9932
8.1255 | -90.9939
8.1258 | -90.9941
8.1262 | -90.9945
8.1265 | -90.9953
8.1268 | -90.9961
8.1271 | -90.9964
8.1274 | -90.9978
8.1277 | -90.9992
8.128 | -91.0002
8.1283 | -91.0007
8.1286 | -91.0012
8.1289 | -91.0015
8.1292 | -91.0012
8.1295 | -91.0015
8.1299 | -91.0025
8.1302 | -91.0032
8.1305 | -91.0038
8.1308 | -91.0048
8.1311 | -91.0047
8.1314 | -91.0048
8.1317 | -91.0056
8.132 | -91.0057
8.1323 | -91.0064
8.1326 | -91.0063
8.1329 | -91.0066
8.1333 | -91.0075
8.1336 | -91.0066
8.1339 | -91.0062
8.1342 | -91.0062
8.1345 | -91.0069
8.1348 | -91.0068
8.1351 | -91.0069
8.1354 | -91.0075
8.1357 | -91.0088
8.136 | -91.0094
8.1363 | -91.011
8.1366 | -91.0113
8.137 | -91.0113
8.1373 | -91.0099
8.1376 | -91.0092
8.1379 | -91.0094
8.1382 | -91.0089
8.1385 | -91.0052
8.1388 | -91.0015
8.1391 | -91.0007
8.1394 | -91.0004
8.1397 | -90.9962
8.14 | -90.9907
8.1404 | -90.9888
8.1407 | -90.9874
8.141 | -90.982
8.1413 | -90.9769
8.1416 | -90.9737
8.1419 | -90.9735
8.1422 | -90.9716
8.1425 | -90.9607
8.1428 | -90.9601
8.1431 | -90.9609
8.1434 | -90.96
8.1438 | -90.958
8.1441 | -90.9568
8.1444 | -90.9575
8.1447 | -90.9556
8.145 | -90.9552
8.1453 | -90.956
8.1456 | -90.9557
8.1459 | -90.9559
8.1462 | -90.9559
8.1465 | -90.9561
8.1468 | -90.9558
8.1471 | -90.9558
8.1475 | -90.9559
8.1478 | -90.9558
8.1481 | -90.955
8.1484 | -90.9545
8.1487 | -90.9542
8.149 | -90.9541
8.1493 | -90.9547
8.1496 | -90.9557
8.1499 | -90.9573
8.1502 | -90.9577
8.1505 | -90.9572
8.1509 | -90.958
8.1512 | -90.9577
8.1515 | -90.9583
8.1518 | -90.9583
8.1521 | -90.9585
8.1524 | -90.9585
8.1527 | -90.9585
8.153 | -90.958
8.1533 | -90.9579
8.1536 | -90.9583
8.1539 | -90.9581
8.1542 | -90.9584
8.1546 | -90.9583
8.1549 | -90.9579
8.1552 | -90.9578
8.1555 | -90.9577
8.1558 | -90.9576
8.1561 | -90.9571
8.1564 | -90.9567
8.1567 | -90.9563
8.157 | -90.956
8.1573 | -90.9559
8.1576 | -90.9554
8.158 | -90.9551
8.1583 | -90.9551
8.1586 | -90.9551
8.1589 | -90.955
8.1592 | -90.9549
8.1595 | -90.9547
8.1598 | -90.954
8.1601 | -90.9523
8.1604 | -90.9517
8.1607 | -90.9515
8.161 | -90.9516
8.1614 | -90.9513
8.1617 | -90.9512
8.162 | -90.951
8.1623 | -90.9513
8.1626 | -90.9514
8.1629 | -90.9511
8.1632 | -90.9509
8.1635 | -90.9512
8.1638 | -90.951
8.1641 | -90.9505
8.1644 | -90.9497
8.1647 | -90.9494
8.1651 | -90.9492
8.1654 | -90.9489
8.1657 | -90.949
8.166 | -90.9488
8.1663 | -90.9487
8.1666 | -90.949
8.1669 | -90.949
8.1672 | -90.9491
8.1675 | -90.9494
8.1678 | -90.9491
8.1681 | -90.9489
8.1685 | -90.9476
8.1688 | -90.9469
8.1691 | -90.946
8.1694 | -90.9465
8.1697 | -90.9462
8.17 | -90.9456
8.1703 | -90.9452
8.1706 | -90.9455
8.1709 | -90.9451
8.1712 | -90.9446
8.1715 | -90.9443
8.1718 | -90.9442
8.1722 | -90.9438
8.1725 | -90.9435
8.1728 | -90.9433
8.1731 | -90.9433
8.1734 | -90.9434
8.1737 | -90.9427
8.174 | -90.9418
8.1743 | -90.9416
8.1746 | -90.9416
8.1749 | -90.941
8.1752 | -90.9408
8.1756 | -90.9404
8.1759 | -90.9401
8.1762 | -90.9404
8.1765 | -90.9401
8.1768 | -90.9397
8.1771 | -90.9393
8.1774 | -90.9388
8.1777 | -90.9377
8.178 | -90.9364
8.1783 | -90.9359
8.1786 | -90.9359
8.1789 | -90.9363
8.1793 | -90.937
8.1796 | -90.9362
8.1799 | -90.934
8.1802 | -90.9332
8.1805 | -90.9326
8.1808 | -90.9314
8.1811 | -90.9312
8.1814 | -90.9316
8.1817 | -90.9309
8.182 | -90.93
8.1823 | -90.9293
8.1827 | -90.9295
8.183 | -90.929
8.1833 | -90.9285
8.1836 | -90.9279
8.1839 | -90.9274
8.1842 | -90.9271
8.1845 | -90.9264
8.1848 | -90.9257
8.1851 | -90.9251
8.1854 | -90.925
8.1857 | -90.9248
8.1861 | -90.9246
8.1864 | -90.924
8.1867 | -90.9234
8.187 | -90.9223
8.1873 | -90.9225
8.1876 | -90.9229
8.1879 | -90.9225
8.1882 | -90.9216
8.1885 | -90.921
8.1888 | -90.9201
8.1891 | -90.919
8.1894 | -90.9183
8.1898 | -90.9179
8.1901 | -90.9174
8.1904 | -90.9161
8.1907 | -90.9153
8.191 | -90.9147
8.1913 | -90.9144
8.1916 | -90.9137
8.1919 | -90.9127
8.1922 | -90.9119
8.1925 | -90.9113
8.1928 | -90.9104
8.1932 | -90.9085
8.1935 | -90.9067
8.1938 | -90.906
8.1941 | -90.9052
8.1944 | -90.9043
8.1947 | -90.9038
8.195 | -90.9037
8.1953 | -90.9023
8.1956 | -90.9014
8.1959 | -90.9001
8.1962 | -90.8987
8.1965 | -90.8969
8.1969 | -90.896
8.1972 | -90.8952
8.1975 | -90.8942
8.1978 | -90.8934
8.1981 | -90.8925
8.1984 | -90.8916
8.1987 | -90.8914
8.199 | -90.8905
8.1993 | -90.8898
8.1996 | -90.889
8.1999 | -90.8882
8.2003 | -90.8881
8.2006 | -90.8874
8.2009 | -90.8851
8.2012 | -90.8839
8.2015 | -90.8829
8.2018 | -90.8817
8.2021 | -90.8802
8.2024 | -90.8784
8.2027 | -90.8769
8.203 | -90.876
8.2033 | -90.8745
8.2037 | -90.8738
8.204 | -90.8729
8.2043 | -90.8713
8.2046 | -90.8704
8.2049 | -90.8685
8.2052 | -90.8673
8.2055 | -90.866
8.2058 | -90.8647
8.2061 | -90.863
8.2064 | -90.8613
8.2067 | -90.8599
8.207 | -90.8589
8.2074 | -90.8573
8.2077 | -90.8562
8.208 | -90.8558
8.2083 | -90.8558
8.2086 | -90.8552
8.2089 | -90.8538
8.2092 | -90.8526
8.2095 | -90.8515
8.2098 | -90.8504
8.2101 | -90.8495
8.2104 | -90.8482
8.2108 | -90.848
8.2111 | -90.8471
8.2114 | -90.8458
8.2117 | -90.8445
8.212 | -90.8436
8.2123 | -90.8428
8.2126 | -90.8416
8.2129 | -90.8405
8.2132 | -90.8391
8.2135 | -90.8372
8.2138 | -90.836
8.2141 | -90.8346
8.2145 | -90.8331
8.2148 | -90.832
8.2151 | -90.8309
8.2154 | -90.8292
8.2157 | -90.827
8.216 | -90.8253
8.2163 | -90.8237
8.2166 | -90.8229
8.2169 | -90.8221
8.2172 | -90.8209
8.2175 | -90.8195
8.2179 | -90.8177
8.2182 | -90.8169
8.2185 | -90.8163
8.2188 | -90.8151
8.2191 | -90.8135
8.2194 | -90.8117
8.2197 | -90.8106
8.22 | -90.8094
8.2203 | -90.8075
8.2206 | -90.8062
8.2209 | -90.8048
8.2213 | -90.8034
8.2216 | -90.802
8.2219 | -90.8007
8.2222 | -90.7994
8.2225 | -90.7973
8.2228 | -90.7965
8.2231 | -90.7952
8.2234 | -90.7938
8.2237 | -90.7923
8.224 | -90.7909
8.2243 | -90.7893
8.2246 | -90.7879
8.225 | -90.7875
8.2253 | -90.7864
8.2256 | -90.7849
8.2259 | -90.7835
8.2262 | -90.7821
8.2265 | -90.7807
8.2268 | -90.7793
8.2271 | -90.778
8.2274 | -90.7772
8.2277 | -90.7764
8.228 | -90.775
8.2284 | -90.7738
8.2287 | -90.7726
8.229 | -90.7712
8.2293 | -90.7697
8.2296 | -90.7687
8.2299 | -90.7672
8.2302 | -90.7648
8.2305 | -90.7626
8.2308 | -90.7615
8.2311 | -90.7607
8.2314 | -90.7582
8.2317 | -90.7569
8.2321 | -90.7554
8.2324 | -90.7539
8.2327 | -90.7532
8.233 | -90.7524
8.2333 | -90.7505
8.2336 | -90.7489
8.2339 | -90.7474
8.2342 | -90.7467
8.2345 | -90.7461
8.2348 | -90.7447
8.2351 | -90.7432
8.2355 | -90.7413
8.2358 | -90.7399
8.2361 | -90.7383
8.2364 | -90.7367
8.2367 | -90.7351
8.237 | -90.7333
8.2373 | -90.7317
8.2376 | -90.73
8.2379 | -90.7283
8.2382 | -90.7268
8.2385 | -90.7254
8.2388 | -90.7241
8.2392 | -90.7227
8.2395 | -90.7215
8.2398 | -90.7203
8.2401 | -90.7191
8.2404 | -90.7164
8.2407 | -90.7147
8.241 | -90.7129
8.2413 | -90.7099
8.2416 | -90.7076
8.2419 | -90.7066
8.2422 | -90.7054
8.2426 | -90.7036
8.2429 | -90.7019
8.2432 | -90.7003
8.2435 | -90.6986
8.2438 | -90.6974
8.2441 | -90.6962
8.2444 | -90.6948
8.2447 | -90.6936
8.245 | -90.6918
8.2453 | -90.69
8.2456 | -90.688
8.246 | -90.6863
8.2463 | -90.684
8.2466 | -90.682
8.2469 | -90.6801
8.2472 | -90.678
8.2475 | -90.6764
8.2478 | -90.6745
8.2481 | -90.6729
8.2484 | -90.6707
8.2487 | -90.6684
8.249 | -90.6667
8.2493 | -90.6646
8.2497 | -90.6629
8.25 | -90.6613
8.2503 | -90.6601
8.2506 | -90.659
8.2509 | -90.6574
8.2512 | -90.6557
8.2515 | -90.6545
8.2518 | -90.6535
8.2521 | -90.6521
8.2524 | -90.6504
8.2527 | -90.6485
8.2531 | -90.6455
8.2534 | -90.644
8.2537 | -90.6428
8.254 | -90.6415
8.2543 | -90.6392
8.2546 | -90.6373
8.2549 | -90.6355
8.2552 | -90.6335
8.2555 | -90.6317
8.2558 | -90.63
8.2561 | -90.6281
8.2564 | -90.6266
8.2568 | -90.6249
8.2571 | -90.6231
8.2574 | -90.6213
8.2577 | -90.6193
8.258 | -90.6175
8.2583 | -90.6152
8.2586 | -90.6137
8.2589 | -90.6117
8.2592 | -90.6089
8.2595 | -90.6067
8.2598 | -90.6048
8.2602 | -90.6028
8.2605 | -90.6013
8.2608 | -90.5995
8.2611 | -90.5977
8.2614 | -90.5957
8.2617 | -90.5938
8.262 | -90.592
8.2623 | -90.5903
8.2626 | -90.5886
8.2629 | -90.5866
8.2632 | -90.5849
8.2636 | -90.5829
8.2639 | -90.5809
8.2642 | -90.5791
8.2645 | -90.5774
8.2648 | -90.575
8.2651 | -90.5731
8.2654 | -90.5719
8.2657 | -90.5712
8.266 | -90.5704
8.2663 | -90.5685
8.2666 | -90.5662
8.2669 | -90.5641
8.2673 | -90.5621
8.2676 | -90.5602
8.2679 | -90.5584
8.2682 | -90.5566
8.2685 | -90.5551
8.2688 | -90.553
8.2691 | -90.5508
8.2694 | -90.549
8.2697 | -90.5475
8.27 | -90.5455
8.2703 | -90.5437
8.2707 | -90.5417
8.271 | -90.5401
8.2713 | -90.5382
8.2716 | -90.5363
8.2719 | -90.5343
8.2722 | -90.5326
8.2725 | -90.5308
8.2728 | -90.5289
8.2731 | -90.527
8.2734 | -90.5251
8.2737 | -90.5237
8.2741 | -90.5225
8.2744 | -90.5209
8.2747 | -90.5192
8.275 | -90.5173
8.2753 | -90.5155
8.2756 | -90.5142
8.2759 | -90.5116
8.2762 | -90.51
8.2765 | -90.5079
8.2768 | -90.506
8.2771 | -90.5037
8.2774 | -90.5017
8.2778 | -90.4996
8.2781 | -90.4978
8.2784 | -90.4957
8.2787 | -90.4937
8.279 | -90.4921
8.2793 | -90.4907
8.2796 | -90.4885
8.2799 | -90.4868
8.2802 | -90.4853
8.2805 | -90.4834
8.2808 | -90.4819
8.2812 | -90.48
8.2815 | -90.4783
8.2818 | -90.4768
8.2821 | -90.4741
8.2824 | -90.4724
8.2827 | -90.4708
8.283 | -90.4698
8.2833 | -90.4688
8.2836 | -90.467
8.2839 | -90.4649
8.2842 | -90.4628
8.2845 | -90.461
8.2849 | -90.4577
8.2852 | -90.455
8.2855 | -90.4529
8.2858 | -90.4508
8.2861 | -90.4492
8.2864 | -90.4472
8.2867 | -90.4454
8.287 | -90.4434
8.2873 | -90.4414
8.2876 | -90.4394
8.2879 | -90.4374
8.2883 | -90.4355
8.2886 | -90.4337
8.2889 | -90.4321
8.2892 | -90.43
8.2895 | -90.428
8.2898 | -90.4261
8.2901 | -90.4241
8.2904 | -90.4221
8.2907 | -90.4201
8.291 | -90.4181
8.2913 | -90.4162
8.2917 | -90.4143
8.292 | -90.4123
8.2923 | -90.4104
8.2926 | -90.4084
8.2929 | -90.4068
8.2932 | -90.4042
8.2935 | -90.4015
8.2938 | -90.3995
8.2941 | -90.3974
8.2944 | -90.3955
8.2947 | -90.3934
8.295 | -90.3916
8.2954 | -90.3896
8.2957 | -90.3876
8.296 | -90.3851
8.2963 | -90.383
8.2966 | -90.3811
8.2969 | -90.3788
8.2972 | -90.3767
8.2975 | -90.3747
8.2978 | -90.3734
8.2981 | -90.3719
8.2984 | -90.37
8.2988 | -90.3683
8.2991 | -90.3664
8.2994 | -90.3645
8.2997 | -90.3625
8.3 | -90.3607
8.3003 | -90.359
8.3006 | -90.3572
8.3009 | -90.3552
8.3012 | -90.3533
8.3015 | -90.3512
8.3018 | -90.3493
8.3021 | -90.3478
8.3025 | -90.346
8.3028 | -90.3442
8.3031 | -90.3435
8.3034 | -90.3418
8.3037 | -90.3403
8.304 | -90.3385
8.3043 | -90.3367
8.3046 | -90.3348
8.3049 | -90.3331
8.3052 | -90.331
8.3055 | -90.3289
8.3059 | -90.3273
8.3062 | -90.3248
8.3065 | -90.3219
8.3068 | -90.3202
8.3071 | -90.3185
8.3074 | -90.317
8.3077 | -90.3157
8.308 | -90.3137
8.3083 | -90.3121
8.3086 | -90.3105
8.3089 | -90.3083
8.3092 | -90.3059
8.3096 | -90.3042
8.3099 | -90.3027
8.3102 | -90.3008
8.3105 | -90.2987
8.3108 | -90.2971
8.3111 | -90.295
8.3114 | -90.2936
8.3117 | -90.2919
8.312 | -90.2902
8.3123 | -90.2886
8.3126 | -90.2868
8.313 | -90.285
8.3133 | -90.2827
8.3136 | -90.2804
8.3139 | -90.2783
8.3142 | -90.2765
8.3145 | -90.2747
8.3148 | -90.2728
8.3151 | -90.2713
8.3154 | -90.2697
8.3157 | -90.2679
8.316 | -90.2663
8.3164 | -90.2656
8.3167 | -90.2648
8.317 | -90.2628
8.3173 | -90.2612
8.3176 | -90.2594
8.3179 | -90.2577
8.3182 | -90.2559
8.3185 | -90.254
8.3188 | -90.2523
8.3191 | -90.2508
8.3194 | -90.2494
8.3197 | -90.2486
8.3201 | -90.2472
8.3204 | -90.2454
8.3207 | -90.2439
8.321 | -90.2422
8.3213 | -90.2408
8.3216 | -90.2391
8.3219 | -90.2372
8.3222 | -90.2354
8.3225 | -90.2335
8.3228 | -90.2318
8.3231 | -90.2297
8.3235 | -90.2278
8.3238 | -90.2261
8.3241 | -90.2242
8.3244 | -90.2227
8.3247 | -90.2212
8.325 | -90.2193
8.3253 | -90.217
8.3256 | -90.2148
8.3259 | -90.2129
8.3262 | -90.2111
8.3265 | -90.2093
8.3268 | -90.2074
8.3272 | -90.2058
8.3275 | -90.2046
8.3278 | -90.2032
8.3281 | -90.202
8.3284 | -90.2009
8.3287 | -90.1994
8.329 | -90.1981
8.3293 | -90.1964
8.3296 | -90.1951
8.3299 | -90.1932
8.3302 | -90.1914
8.3306 | -90.1896
8.3309 | -90.1875
8.3312 | -90.1855
8.3315 | -90.1839
8.3318 | -90.1821
8.3321 | -90.1802
8.3324 | -90.1784
8.3327 | -90.1764
8.333 | -90.1741
8.3333 | -90.1723
8.3336 | -90.1704
8.334 | -90.1687
8.3343 | -90.1672
8.3346 | -90.1655
8.3349 | -90.1639
8.3352 | -90.1626
8.3355 | -90.1612
8.3358 | -90.1595
8.3361 | -90.1578
8.3364 | -90.1561
8.3367 | -90.1547
8.337 | -90.1533
8.3373 | -90.1513
8.3377 | -90.1495
8.338 | -90.1479
8.3383 | -90.147
8.3386 | -90.1452
8.3389 | -90.1434
8.3392 | -90.1416
8.3395 | -90.1401
8.3398 | -90.1388
8.3401 | -90.1375
8.3404 | -90.1358
8.3407 | -90.1345
8.3411 | -90.1332
8.3414 | -90.1315
8.3417 | -90.1297
8.342 | -90.1281
8.3423 | -90.1263
8.3426 | -90.1244
8.3429 | -90.1223
8.3432 | -90.1203
8.3435 | -90.119
8.3438 | -90.1174
8.3441 | -90.1162
8.3444 | -90.1149
8.3448 | -90.1135
8.3451 | -90.1119
8.3454 | -90.1102
8.3457 | -90.109
8.346 | -90.107
8.3463 | -90.1053
8.3466 | -90.1038
8.3469 | -90.1022
8.3472 | -90.1016
8.3475 | -90.1005
8.3478 | -90.0991
8.3482 | -90.0975
8.3485 | -90.0954
8.3488 | -90.0936
8.3491 | -90.0917
8.3494 | -90.0904
8.3497 | -90.0887
8.35 | -90.0869
8.3503 | -90.0852
8.3506 | -90.0838
8.3509 | -90.0822
8.3512 | -90.0808
8.3516 | -90.0796
8.3519 | -90.0782
8.3522 | -90.0766
8.3525 | -90.0752
8.3528 | -90.0738
8.3531 | -90.0724
8.3534 | -90.0708
8.3537 | -90.0693
8.354 | -90.0681
8.3543 | -90.0669
8.3546 | -90.0654
8.3549 | -90.0641
8.3553 | -90.0635
8.3556 | -90.0623
8.3559 | -90.0609
8.3562 | -90.0596
8.3565 | -90.0574
8.3568 | -90.0554
8.3571 | -90.0539
8.3574 | -90.0526
8.3577 | -90.0515
8.358 | -90.0501
8.3583 | -90.0489
8.3587 | -90.0479
8.359 | -90.0465
8.3593 | -90.0452
8.3596 | -90.044
8.3599 | -90.0427
8.3602 | -90.0412
8.3605 | -90.0398
8.3608 | -90.0377
8.3611 | -90.0356
8.3614 | -90.0341
8.3617 | -90.0328
8.362 | -90.0316
8.3624 | -90.0299
8.3627 | -90.0286
8.363 | -90.028
8.3633 | -90.0267
8.3636 | -90.0255
8.3639 | -90.0244
8.3642 | -90.023
8.3645 | -90.0219
8.3648 | -90.0214
8.3651 | -90.0214
8.3654 | -90.021
8.3658 | -90.0204
8.3661 | -90.0194
8.3664 | -90.0181
8.3667 | -90.0167
8.367 | -90.0157
8.3673 | -90.0147
8.3676 | -90.0134
8.3679 | -90.012
8.3682 | -90.0108
8.3685 | -90.0098
8.3688 | -90.009
8.3691 | -90.0077
8.3695 | -90.0065
8.3698 | -90.0054
8.3701 | -90.0044
8.3704 | -90.0036
8.3707 | -90.0018
8.371 | -89.9993
8.3713 | -89.9977
8.3716 | -89.9964
8.3719 | -89.995
8.3722 | -89.9939
8.3725 | -89.9927
8.3729 | -89.9917
8.3732 | -89.9914
8.3735 | -89.9897
8.3738 | -89.9875
8.3741 | -89.9874
8.3744 | -89.9862
8.3747 | -89.9853
8.375 | -89.9843
8.3753 | -89.9832
8.3756 | -89.9826
8.3759 | -89.982
8.3763 | -89.981
8.3766 | -89.9797
8.3769 | -89.9785
8.3772 | -89.9775
8.3775 | -89.9764
8.3778 | -89.9755
8.3781 | -89.9743
8.3784 | -89.9733
8.3787 | -89.9721
8.379 | -89.9712
8.3793 | -89.9712
8.3796 | -89.9714
8.38 | -89.9711
8.3803 | -89.9703
8.3806 | -89.969
8.3809 | -89.9678
8.3812 | -89.9667
8.3815 | -89.9659
8.3818 | -89.9646
8.3821 | -89.9634
8.3824 | -89.9625
8.3827 | -89.9614
8.383 | -89.9619
8.3834 | -89.9613
8.3837 | -89.9607
8.384 | -89.9598
8.3843 | -89.9591
8.3846 | -89.9583
8.3849 | -89.9575
8.3852 | -89.9566
8.3855 | -89.9559
8.3858 | -89.9551
8.3861 | -89.9537
8.3864 | -89.9523
8.3867 | -89.9515
8.3871 | -89.9507
8.3874 | -89.9501
8.3877 | -89.9493
8.388 | -89.9486
8.3883 | -89.9478
8.3886 | -89.9469
8.3889 | -89.9461
8.3892 | -89.9452
8.3895 | -89.9443
8.3898 | -89.9435
8.3901 | -89.9426
8.3905 | -89.9419
8.3908 | -89.9407
8.3911 | -89.9397
8.3914 | -89.9391
8.3917 | -89.9384
8.392 | -89.9388
8.3923 | -89.9377
8.3926 | -89.9369
8.3929 | -89.9366
8.3932 | -89.936
8.3935 | -89.9354
8.3939 | -89.9347
8.3942 | -89.9338
8.3945 | -89.9338
8.3948 | -89.934
8.3951 | -89.9334
8.3954 | -89.9328
8.3957 | -89.932
8.396 | -89.9315
8.3963 | -89.9307
8.3966 | -89.9299
8.3969 | -89.929
8.3972 | -89.9284
8.3976 | -89.9275
8.3979 | -89.9266
8.3982 | -89.9259
8.3985 | -89.9253
8.3988 | -89.9248
8.3991 | -89.9241
8.3994 | -89.923
8.3997 | -89.9223
8.4 | -89.9217
8.4003 | -89.9211
8.4006 | -89.9216
8.401 | -89.9209
8.4013 | -89.9205
8.4016 | -89.9201
8.4019 | -89.9197
8.4022 | -89.9189
8.4025 | -89.9181
8.4028 | -89.9174
8.4031 | -89.9169
8.4034 | -89.9165
8.4037 | -89.916
8.404 | -89.9155
8.4044 | -89.915
8.4047 | -89.9146
8.405 | -89.9142
8.4053 | -89.9137
8.4056 | -89.9133
8.4059 | -89.9127
8.4062 | -89.912
8.4065 | -89.9115
8.4068 | -89.9111
8.4071 | -89.9104
8.4074 | -89.9095
8.4077 | -89.9086
8.4081 | -89.9081
8.4084 | -89.9079
8.4087 | -89.9072
8.409 | -89.9067
8.4093 | -89.9064
8.4096 | -89.9074
8.4099 | -89.907
8.4102 | -89.9066
8.4105 | -89.9063
8.4108 | -89.906
8.4111 | -89.9057
8.4115 | -89.9056
8.4118 | -89.9055
8.4121 | -89.9049
8.4124 | -89.9046
8.4127 | -89.9043
8.413 | -89.9047
8.4133 | -89.9046
8.4136 | -89.9043
8.4139 | -89.9039
8.4142 | -89.9037
8.4145 | -89.9036
8.4148 | -89.9034
8.4152 | -89.9035
8.4155 | -89.9034
8.4158 | -89.9027
8.4161 | -89.9019
8.4164 | -89.9018
8.4167 | -89.9017
8.417 | -89.9012
8.4173 | -89.9011
8.4176 | -89.9008
8.4179 | -89.9005
8.4182 | -89.9002
8.4186 | -89.8991
8.4189 | -89.8996
8.4192 | -89.8998
8.4195 | -89.8997
8.4198 | -89.8998
8.4201 | -89.8996
8.4204 | -89.8996
8.4207 | -89.8997
8.421 | -89.8991
8.4213 | -89.8989
8.4216 | -89.8988
8.422 | -89.8986
8.4223 | -89.8987
8.4226 | -89.8979
8.4229 | -89.8972
8.4232 | -89.8971
8.4235 | -89.8971
8.4238 | -89.897
8.4241 | -89.8968
8.4244 | -89.8965
8.4247 | -89.8963
8.425 | -89.8962
8.4253 | -89.8962
8.4257 | -89.8963
8.426 | -89.8963
8.4263 | -89.8964
8.4266 | -89.8964
8.4269 | -89.8965
8.4272 | -89.8964
8.4275 | -89.8954
8.4278 | -89.8943
8.4281 | -89.8945
8.4284 | -89.8948
8.4287 | -89.8952
8.4291 | -89.8956
8.4294 | -89.8954
8.4297 | -89.8949
8.43 | -89.894
8.4303 | -89.8934
8.4306 | -89.893
8.4309 | -89.8934
8.4312 | -89.8933
8.4315 | -89.8936
8.4318 | -89.8939
8.4321 | -89.8939
8.4324 | -89.8937
8.4328 | -89.8934
8.4331 | -89.8932
8.4334 | -89.893
8.4337 | -89.8933
8.434 | -89.8933
8.4343 | -89.8932
8.4346 | -89.8932
8.4349 | -89.8935
8.4352 | -89.8937
8.4355 | -89.8941
8.4358 | -89.894
8.4362 | -89.8939
8.4365 | -89.8947
8.4368 | -89.8944
8.4371 | -89.8944
8.4374 | -89.8945
8.4377 | -89.8949
8.438 | -89.8951
8.4383 | -89.8951
8.4386 | -89.8951
8.4389 | -89.8953
8.4392 | -89.8954
8.4395 | -89.8957
8.4399 | -89.8959
8.4402 | -89.8961
8.4405 | -89.8962
8.4408 | -89.8966
8.4411 | -89.8968
8.4414 | -89.8967
8.4417 | -89.8971
8.442 | -89.8975
8.4423 | -89.8982
8.4426 | -89.8987
8.4429 | -89.899
8.4433 | -89.8991
8.4436 | -89.8995
8.4439 | -89.8999
8.4442 | -89.9001
8.4445 | -89.8999
8.4448 | -89.8995
8.4451 | -89.9
8.4454 | -89.9016
8.4457 | -89.902
8.446 | -89.9022
8.4463 | -89.9027
8.4467 | -89.9032
8.447 | -89.9037
8.4473 | -89.9044
8.4476 | -89.905
8.4479 | -89.9053
8.4482 | -89.906
8.4485 | -89.9066
8.4488 | -89.9069
8.4491 | -89.9074
8.4494 | -89.9084
8.4497 | -89.9093
8.45 | -89.9101
8.4504 | -89.9109
8.4507 | -89.9114
8.451 | -89.9119
8.4513 | -89.9124
8.4516 | -89.9129
8.4519 | -89.9137
8.4522 | -89.9142
8.4525 | -89.9149
8.4528 | -89.9156
8.4531 | -89.9161
8.4534 | -89.9162
8.4538 | -89.9161
8.4541 | -89.9173
8.4544 | -89.9181
8.4547 | -89.9189
8.455 | -89.9195
8.4553 | -89.9199
8.4556 | -89.9204
8.4559 | -89.9214
8.4562 | -89.9228
8.4565 | -89.9239
8.4568 | -89.9246
8.4571 | -89.9254
8.4575 | -89.926
8.4578 | -89.9263
8.4581 | -89.9267
8.4584 | -89.9272
8.4587 | -89.9279
8.459 | -89.9285
8.4593 | -89.9292
8.4596 | -89.9297
8.4599 | -89.9305
8.4602 | -89.9314
8.4605 | -89.9318
8.4609 | -89.9326
8.4612 | -89.9327
8.4615 | -89.9333
8.4618 | -89.9341
8.4621 | -89.9351
8.4624 | -89.936
8.4627 | -89.9371
8.463 | -89.9389
8.4633 | -89.9406
8.4636 | -89.9418
8.4639 | -89.9425
8.4643 | -89.9434
8.4646 | -89.9442
8.4649 | -89.9448
8.4652 | -89.9456
8.4655 | -89.9465
8.4658 | -89.9473
8.4661 | -89.9482
8.4664 | -89.9491
8.4667 | -89.9501
8.467 | -89.9511
8.4673 | -89.9519
8.4676 | -89.9526
8.468 | -89.9536
8.4683 | -89.9548
8.4686 | -89.9558
8.4689 | -89.9564
8.4692 | -89.9567
8.4695 | -89.9572
8.4698 | -89.9581
8.4701 | -89.9588
8.4704 | -89.9592
8.4707 | -89.9602
8.471 | -89.961
8.4714 | -89.9618
8.4717 | -89.963
8.472 | -89.9647
8.4723 | -89.9655
8.4726 | -89.9665
8.4729 | -89.9674
8.4732 | -89.9684
8.4735 | -89.9696
8.4738 | -89.9705
8.4741 | -89.9713
8.4744 | -89.9729
8.4747 | -89.9747
8.4751 | -89.9756
8.4754 | -89.9764
8.4757 | -89.9775
8.476 | -89.9787
8.4763 | -89.9797
8.4766 | -89.9804
8.4769 | -89.981
8.4772 | -89.9817
8.4775 | -89.9822
8.4778 | -89.9832
8.4781 | -89.984
8.4785 | -89.9841
8.4788 | -89.9849
8.4791 | -89.9861
8.4794 | -89.9872
8.4797 | -89.9882
8.48 | -89.9893
8.4803 | -89.9907
8.4806 | -89.9922
8.4809 | -89.9943
8.4812 | -89.9951
8.4815 | -89.9959
8.4819 | -89.9969
8.4822 | -89.9982
8.4825 | -89.9994
8.4828 | -90.0005
8.4831 | -90.0019
8.4834 | -90.0036
8.4837 | -90.0047
8.484 | -90.0056
8.4843 | -90.0067
8.4846 | -90.0078
8.4849 | -90.0092
8.4852 | -90.0109
8.4856 | -90.0121
8.4859 | -90.0136
8.4862 | -90.0149
8.4865 | -90.0161
8.4868 | -90.0173
8.4871 | -90.0185
8.4874 | -90.0196
8.4877 | -90.0207
8.488 | -90.0223
8.4883 | -90.0245
8.4886 | -90.026
8.489 | -90.0269
8.4893 | -90.0281
8.4896 | -90.0293
8.4899 | -90.0314
8.4902 | -90.033
8.4905 | -90.0341
8.4908 | -90.0351
8.4911 | -90.0362
8.4914 | -90.0374
8.4917 | -90.0382
8.492 | -90.0385
8.4923 | -90.0397
8.4927 | -90.0411
8.493 | -90.0424
8.4933 | -90.0438
8.4936 | -90.0449
8.4939 | -90.0463
8.4942 | -90.0475
8.4945 | -90.0478
8.4948 | -90.0484
8.4951 | -90.0497
8.4954 | -90.0507
8.4957 | -90.0519
8.4961 | -90.0531
8.4964 | -90.0539
8.4967 | -90.0542
8.497 | -90.0557
8.4973 | -90.0573
8.4976 | -90.0585
8.4979 | -90.0598
8.4982 | -90.0612
8.4985 | -90.0622
8.4988 | -90.0642
8.4991 | -90.0654
8.4994 | -90.0665
8.4998 | -90.0682
8.5001 | -90.0703
8.5004 | -90.0719
8.5007 | -90.0733
8.501 | -90.0746
8.5013 | -90.0759
8.5016 | -90.0774
8.5019 | -90.0787
8.5022 | -90.0799
8.5025 | -90.0812
8.5028 | -90.0828
8.5032 | -90.0842
8.5035 | -90.0855
8.5038 | -90.0866
8.5041 | -90.0877
8.5044 | -90.0892
8.5047 | -90.0907
8.505 | -90.0921
8.5053 | -90.0935
8.5056 | -90.0946
8.5059 | -90.0954
8.5062 | -90.0964
8.5066 | -90.0978
8.5069 | -90.0993
8.5072 | -90.1009
8.5075 | -90.1031
8.5078 | -90.1049
8.5081 | -90.1069
8.5084 | -90.1087
8.5087 | -90.11
8.509 | -90.1113
8.5093 | -90.1127
8.5096 | -90.114
8.5099 | -90.1154
8.5103 | -90.1166
8.5106 | -90.1181
8.5109 | -90.1194
8.5112 | -90.1208
8.5115 | -90.1223
8.5118 | -90.1237
8.5121 | -90.1251
8.5124 | -90.1265
8.5127 | -90.1281
8.513 | -90.1297
8.5133 | -90.1314
8.5137 | -90.1328
8.514 | -90.1343
8.5143 | -90.1361
8.5146 | -90.1381
8.5149 | -90.1395
8.5152 | -90.1409
8.5155 | -90.1424
8.5158 | -90.1438
8.5161 | -90.1453
8.5164 | -90.1474
8.5167 | -90.1489
8.517 | -90.1505
8.5174 | -90.1521
8.5177 | -90.1536
8.518 | -90.1551
8.5183 | -90.1565
8.5186 | -90.158
8.5189 | -90.1595
8.5192 | -90.161
8.5195 | -90.1626
8.5198 | -90.1641
8.5201 | -90.1652
8.5204 | -90.1666
8.5208 | -90.1682
8.5211 | -90.1705
8.5214 | -90.1726
8.5217 | -90.1744
8.522 | -90.1762
8.5223 | -90.1777
8.5226 | -90.1789
8.5229 | -90.1803
8.5232 | -90.182
8.5235 | -90.1835
8.5238 | -90.1848
8.5242 | -90.1861
8.5245 | -90.1873
8.5248 | -90.1887
8.5251 | -90.19
8.5254 | -90.192
8.5257 | -90.1936
8.526 | -90.1953
8.5263 | -90.1969
8.5266 | -90.1985
8.5269 | -90.1998
8.5272 | -90.201
8.5275 | -90.2022
8.5279 | -90.2036
8.5282 | -90.2052
8.5285 | -90.2067
8.5288 | -90.2082
8.5291 | -90.21
8.5294 | -90.2116
8.5297 | -90.2134
8.53 | -90.2145
8.5303 | -90.2162
8.5306 | -90.2181
8.5309 | -90.2198
8.5313 | -90.2212
8.5316 | -90.2221
8.5319 | -90.2234
8.5322 | -90.2248
8.5325 | -90.2264
8.5328 | -90.2281
8.5331 | -90.2297
8.5334 | -90.2315
8.5337 | -90.2333
8.534 | -90.2351
8.5343 | -90.2372
8.5346 | -90.2388
8.535 | -90.2402
8.5353 | -90.2417
8.5356 | -90.2431
8.5359 | -90.2448
8.5362 | -90.2466
8.5365 | -90.2483
8.5368 | -90.2502
8.5371 | -90.2521
8.5374 | -90.2539
8.5377 | -90.2555
8.538 | -90.2572
8.5384 | -90.259
8.5387 | -90.2607
8.539 | -90.2621
8.5393 | -90.2636
8.5396 | -90.2649
8.5399 | -90.2661
8.5402 | -90.2672
8.5405 | -90.2686
8.5408 | -90.2702
8.5411 | -90.2719
8.5414 | -90.2738
8.5418 | -90.2754
8.5421 | -90.277
8.5424 | -90.2786
8.5427 | -90.2803
8.543 | -90.282
8.5433 | -90.2838
8.5436 | -90.2854
8.5439 | -90.287
8.5442 | -90.2885
8.5445 | -90.2901
8.5448 | -90.2917
8.5451 | -90.293
8.5455 | -90.2948
8.5458 | -90.2964
8.5461 | -90.2981
8.5464 | -90.2997
8.5467 | -90.3013
8.547 | -90.3027
8.5473 | -90.3043
8.5476 | -90.3063
8.5479 | -90.3085
8.5482 | -90.3102
8.5485 | -90.3124
8.5489 | -90.3146
8.5492 | -90.3162
8.5495 | -90.3174
8.5498 | -90.3188
8.5501 | -90.3203
8.5504 | -90.3218
8.5507 | -90.3235
8.551 | -90.3254
8.5513 | -90.3276
8.5516 | -90.3294
8.5519 | -90.3308
8.5523 | -90.3326
8.5526 | -90.334
8.5529 | -90.335
8.5532 | -90.3366
8.5535 | -90.3383
8.5538 | -90.3401
8.5541 | -90.3419
8.5544 | -90.3439
8.5547 | -90.3455
8.555 | -90.3466
8.5553 | -90.348
8.5556 | -90.3497
8.556 | -90.3512
8.5563 | -90.3517
8.5566 | -90.3527
8.5569 | -90.3543
8.5572 | -90.356
8.5575 | -90.3577
8.5578 | -90.3593
8.5581 | -90.3606
8.5584 | -90.3623
8.5587 | -90.3637
8.559 | -90.3649
8.5594 | -90.3666
8.5597 | -90.368
8.56 | -90.3696
8.5603 | -90.3709
8.5606 | -90.3719
8.5609 | -90.3735
8.5612 | -90.3754
8.5615 | -90.3771
8.5618 | -90.3787
8.5621 | -90.3804
8.5624 | -90.382
8.5627 | -90.3835
8.5631 | -90.3853
8.5634 | -90.3873
8.5637 | -90.3887
8.564 | -90.3899
8.5643 | -90.3911
8.5646 | -90.3932
8.5649 | -90.3955
8.5652 | -90.3972
8.5655 | -90.3989
8.5658 | -90.4005
8.5661 | -90.4022
8.5665 | -90.4037
8.5668 | -90.4048
8.5671 | -90.4062
8.5674 | -90.4079
8.5677 | -90.4095
8.568 | -90.4113
8.5683 | -90.413
8.5686 | -90.4144
8.5689 | -90.4161
8.5692 | -90.4181
8.5695 | -90.4196
8.5698 | -90.4207
8.5702 | -90.4216
8.5705 | -90.4232
8.5708 | -90.4249
8.5711 | -90.4267
8.5714 | -90.428
8.5717 | -90.4293
8.572 | -90.4308
8.5723 | -90.4321
8.5726 | -90.4336
8.5729 | -90.4355
8.5732 | -90.4369
8.5736 | -90.4379
8.5739 | -90.4396
8.5742 | -90.4416
8.5745 | -90.4434
8.5748 | -90.4451
8.5751 | -90.4467
8.5754 | -90.4488
8.5757 | -90.4507
8.576 | -90.4523
8.5763 | -90.4538
8.5766 | -90.4557
8.577 | -90.4574
8.5773 | -90.4592
8.5776 | -90.4612
8.5779 | -90.4627
8.5782 | -90.4643
8.5785 | -90.4659
8.5788 | -90.4671
8.5791 | -90.4685
8.5794 | -90.4698
8.5797 | -90.4717
8.58 | -90.4736
8.5803 | -90.4752
8.5807 | -90.4763
8.581 | -90.4775
8.5813 | -90.4789
8.5816 | -90.4803
8.5819 | -90.4819
8.5822 | -90.4837
8.5825 | -90.4859
8.5828 | -90.4879
8.5831 | -90.4891
8.5834 | -90.4906
8.5837 | -90.4925
8.5841 | -90.4941
8.5844 | -90.4955
8.5847 | -90.4969
8.585 | -90.4976
8.5853 | -90.4992
8.5856 | -90.5005
8.5859 | -90.502
8.5862 | -90.5036
8.5865 | -90.5052
8.5868 | -90.5069
8.5871 | -90.5083
8.5874 | -90.5099
8.5878 | -90.5113
8.5881 | -90.5127
8.5884 | -90.514
8.5887 | -90.5156
8.589 | -90.5172
8.5893 | -90.5189
8.5896 | -90.5202
8.5899 | -90.5217
8.5902 | -90.5238
8.5905 | -90.5262
8.5908 | -90.5281
8.5912 | -90.5291
8.5915 | -90.5305
8.5918 | -90.5321
8.5921 | -90.5337
8.5924 | -90.535
8.5927 | -90.5363
8.593 | -90.5381
8.5933 | -90.5394
8.5936 | -90.5408
8.5939 | -90.5422
8.5942 | -90.5436
8.5946 | -90.5452
8.5949 | -90.5468
8.5952 | -90.5485
8.5955 | -90.5501
8.5958 | -90.5516
8.5961 | -90.5532
8.5964 | -90.5549
8.5967 | -90.5561
8.597 | -90.5576
8.5973 | -90.5594
8.5976 | -90.5615
8.5979 | -90.563
8.5983 | -90.5646
8.5986 | -90.5663
8.5989 | -90.5677
8.5992 | -90.5692
8.5995 | -90.5707
8.5998 | -90.5722
8.6001 | -90.5738
8.6004 | -90.5751
8.6007 | -90.5762
8.601 | -90.5774
8.6013 | -90.5788
8.6017 | -90.5808
8.602 | -90.5825
8.6023 | -90.584
8.6026 | -90.5862
8.6029 | -90.5882
8.6032 | -90.5895
8.6035 | -90.5912
8.6038 | -90.5929
8.6041 | -90.5943
8.6044 | -90.5956
8.6047 | -90.597
8.605 | -90.5983
8.6054 | -90.5999
8.6057 | -90.6007
8.606 | -90.603
8.6063 | -90.6054
8.6066 | -90.6073
8.6069 | -90.6087
8.6072 | -90.6102
8.6075 | -90.6116
8.6078 | -90.613
8.6081 | -90.6145
8.6084 | -90.6161
8.6088 | -90.6176
8.6091 | -90.619
8.6094 | -90.6204
8.6097 | -90.6221
8.61 | -90.6232
8.6103 | -90.6243
8.6106 | -90.6259
8.6109 | -90.6271
8.6112 | -90.6287
8.6115 | -90.6305
8.6118 | -90.6321
8.6122 | -90.6337
8.6125 | -90.6351
8.6128 | -90.6365
8.6131 | -90.638
8.6134 | -90.6397
8.6137 | -90.6411
8.614 | -90.6426
8.6143 | -90.6442
8.6146 | -90.6455
8.6149 | -90.6467
8.6152 | -90.6479
8.6155 | -90.6494
8.6159 | -90.6508
8.6162 | -90.6522
8.6165 | -90.6537
8.6168 | -90.6551
8.6171 | -90.6564
8.6174 | -90.6578
8.6177 | -90.6593
8.618 | -90.6605
8.6183 | -90.6627
8.6186 | -90.6648
8.6189 | -90.6662
8.6193 | -90.6675
8.6196 | -90.6688
8.6199 | -90.6706
8.6202 | -90.6727
8.6205 | -90.6751
8.6208 | -90.6772
8.6211 | -90.679
8.6214 | -90.6804
8.6217 | -90.6817
8.622 | -90.6833
8.6223 | -90.6851
8.6226 | -90.687
8.623 | -90.688
8.6233 | -90.6897
8.6236 | -90.6904
8.6239 | -90.6919
8.6242 | -90.6938
8.6245 | -90.6966
8.6248 | -90.6997
8.6251 | -90.7011
8.6254 | -90.7024
8.6257 | -90.7035
8.626 | -90.7047
8.6264 | -90.706
8.6267 | -90.7077
8.627 | -90.709
8.6273 | -90.7103
8.6276 | -90.7118
8.6279 | -90.713
8.6282 | -90.7147
8.6285 | -90.716
8.6288 | -90.7171
8.6291 | -90.7183
8.6294 | -90.7199
8.6297 | -90.7211
8.6301 | -90.7219
8.6304 | -90.7229
8.6307 | -90.7244
8.631 | -90.7267
8.6313 | -90.7289
8.6316 | -90.7302
8.6319 | -90.7319
8.6322 | -90.7341
8.6325 | -90.7349
8.6328 | -90.7365
8.6331 | -90.7379
8.6335 | -90.7391
8.6338 | -90.7405
8.6341 | -90.7417
8.6344 | -90.7427
8.6347 | -90.7438
8.635 | -90.7451
8.6353 | -90.7462
8.6356 | -90.7474
8.6359 | -90.7485
8.6362 | -90.7498
8.6365 | -90.7517
8.6369 | -90.7536
8.6372 | -90.7554
8.6375 | -90.7567
8.6378 | -90.7579
8.6381 | -90.759
8.6384 | -90.7603
8.6387 | -90.7617
8.639 | -90.763
8.6393 | -90.765
8.6396 | -90.7663
8.6399 | -90.7676
8.6402 | -90.7688
8.6406 | -90.7703
8.6409 | -90.772
8.6412 | -90.7728
8.6415 | -90.7739
8.6418 | -90.7746
8.6421 | -90.7756
8.6424 | -90.7768
8.6427 | -90.778
8.643 | -90.7796
8.6433 | -90.7809
8.6436 | -90.7822
8.644 | -90.7833
8.6443 | -90.7846
8.6446 | -90.7861
8.6449 | -90.7877
8.6452 | -90.7883
8.6455 | -90.7899
8.6458 | -90.7911
8.6461 | -90.7923
8.6464 | -90.7935
8.6467 | -90.7949
8.647 | -90.796
8.6473 | -90.7972
8.6477 | -90.7988
8.648 | -90.8003
8.6483 | -90.802
8.6486 | -90.8034
8.6489 | -90.8046
8.6492 | -90.8056
8.6495 | -90.8073
8.6498 | -90.8086
8.6501 | -90.8092
8.6504 | -90.8104
8.6507 | -90.8118
8.6511 | -90.8131
8.6514 | -90.8141
8.6517 | -90.8152
8.652 | -90.8165
8.6523 | -90.8182
8.6526 | -90.8197
8.6529 | -90.8212
8.6532 | -90.8228
8.6535 | -90.8249
8.6538 | -90.8256
8.6541 | -90.828
8.6545 | -90.8299
8.6548 | -90.831
8.6551 | -90.8321
8.6554 | -90.8334
8.6557 | -90.8345
8.656 | -90.8354
8.6563 | -90.8365
8.6566 | -90.8377
8.6569 | -90.8397
8.6572 | -90.8414
8.6575 | -90.8423
8.6578 | -90.8422
8.6582 | -90.8428
8.6585 | -90.844
8.6588 | -90.8452
8.6591 | -90.846
8.6594 | -90.8468
8.6597 | -90.848
8.66 | -90.8492
8.6603 | -90.8499
8.6606 | -90.8514
8.6609 | -90.8527
8.6612 | -90.8537
8.6616 | -90.8544
8.6619 | -90.8551
8.6622 | -90.8569
8.6625 | -90.858
8.6628 | -90.8589
8.6631 | -90.8599
8.6634 | -90.8609
8.6637 | -90.8621
8.664 | -90.8636
8.6643 | -90.865
8.6646 | -90.8663
8.6649 | -90.8674
8.6653 | -90.8684
8.6656 | -90.8696
8.6659 | -90.8708
8.6662 | -90.8715
8.6665 | -90.8725
8.6668 | -90.8737
8.6671 | -90.8751
8.6674 | -90.8765
8.6677 | -90.8777
8.668 | -90.8783
8.6683 | -90.8787
8.6687 | -90.8789
8.669 | -90.8792
8.6693 | -90.8798
8.6696 | -90.8808
8.6699 | -90.8817
8.6702 | -90.8824
8.6705 | -90.8829
8.6708 | -90.8841
8.6711 | -90.8852
8.6714 | -90.8863
8.6717 | -90.8874
8.6721 | -90.8883
8.6724 | -90.8893
8.6727 | -90.8904
8.673 | -90.8914
8.6733 | -90.8925
8.6736 | -90.8936
8.6739 | -90.8945
8.6742 | -90.8954
8.6745 | -90.8965
8.6748 | -90.8971
8.6751 | -90.8978
8.6754 | -90.899
8.6758 | -90.9003
8.6761 | -90.9013
8.6764 | -90.9023
8.6767 | -90.9035
8.677 | -90.9042
8.6773 | -90.9057
8.6776 | -90.9067
8.6779 | -90.9079
8.6782 | -90.9091
8.6785 | -90.9101
8.6788 | -90.9116
8.6792 | -90.9131
8.6795 | -90.9143
8.6798 | -90.9156
8.6801 | -90.9166
8.6804 | -90.9179
8.6807 | -90.9194
8.681 | -90.9204
8.6813 | -90.9212
8.6816 | -90.9221
8.6819 | -90.9229
8.6822 | -90.9238
8.6826 | -90.9247
8.6829 | -90.9258
8.6832 | -90.9273
8.6835 | -90.9284
8.6838 | -90.9293
8.6841 | -90.9303
8.6844 | -90.9314
8.6847 | -90.9326
8.685 | -90.9338
8.6853 | -90.9347
8.6856 | -90.9353
8.6859 | -90.9363
8.6863 | -90.9375
8.6866 | -90.9388
8.6869 | -90.9398
8.6872 | -90.9411
8.6875 | -90.9422
8.6878 | -90.944
8.6881 | -90.9459
8.6884 | -90.9475
8.6887 | -90.9488
8.689 | -90.9498
8.6893 | -90.951
8.6897 | -90.952
8.69 | -90.953
8.6903 | -90.9542
8.6906 | -90.9553
8.6909 | -90.9562
8.6912 | -90.9573
8.6915 | -90.9586
8.6918 | -90.9588
8.6921 | -90.9591
8.6924 | -90.9606
8.6927 | -90.9617
8.693 | -90.9628
8.6934 | -90.9642
8.6937 | -90.9651
8.694 | -90.9662
8.6943 | -90.9679
8.6946 | -90.9691
8.6949 | -90.9701
8.6952 | -90.971
8.6955 | -90.9721
8.6958 | -90.9733
8.6961 | -90.9744
8.6964 | -90.9761
8.6968 | -90.9779
8.6971 | -90.9785
8.6974 | -90.9793
8.6977 | -90.9804
8.698 | -90.9818
8.6983 | -90.9829
8.6986 | -90.9837
8.6989 | -90.9848
8.6992 | -90.9858
8.6995 | -90.9869
8.6998 | -90.9877
8.7001 | -90.9881
8.7005 | -90.9886
8.7008 | -90.9894
8.7011 | -90.9907
8.7014 | -90.9918
8.7017 | -90.9924
8.702 | -90.9931
8.7023 | -90.9941
8.7026 | -90.995
8.7029 | -90.9962
8.7032 | -90.9969
8.7035 | -90.9972
8.7039 | -90.9978
8.7042 | -90.9988
8.7045 | -91
8.7048 | -91.0007
8.7051 | -91.002
8.7054 | -91.0037
8.7057 | -91.0048
8.706 | -91.006
8.7063 | -91.0069
8.7066 | -91.0078
8.7069 | -91.0085
8.7073 | -91.0087
8.7076 | -91.0091
8.7079 | -91.0101
8.7082 | -91.0107
8.7085 | -91.0105
8.7088 | -91.0108
8.7091 | -91.0113
8.7094 | -91.0122
8.7097 | -91.0133
8.71 | -91.0148
8.7103 | -91.0158
8.7106 | -91.0167
8.711 | -91.0176
8.7113 | -91.0189
8.7116 | -91.02
8.7119 | -91.0209
8.7122 | -91.0218
8.7125 | -91.023
8.7128 | -91.024
8.7131 | -91.025
8.7134 | -91.0269
8.7137 | -91.0294
8.714 | -91.0307
8.7144 | -91.0313
8.7147 | -91.0319
8.715 | -91.0329
8.7153 | -91.0339
8.7156 | -91.0353
8.7159 | -91.0365
8.7162 | -91.0376
8.7165 | -91.039
8.7168 | -91.0398
8.7171 | -91.0416
8.7174 | -91.0437
8.7177 | -91.0445
8.7181 | -91.0462
8.7184 | -91.0477
8.7187 | -91.0497
8.719 | -91.0512
8.7193 | -91.052
8.7196 | -91.0539
8.7199 | -91.0555
8.7202 | -91.0569
8.7205 | -91.058
8.7208 | -91.0594
8.7211 | -91.0603
8.7215 | -91.0615
8.7218 | -91.0633
8.7221 | -91.0647
8.7224 | -91.0674
8.7227 | -91.0694
8.723 | -91.0706
8.7233 | -91.0727
8.7236 | -91.0775
8.7239 | -91.0824
8.7242 | -91.0844
8.7245 | -91.0862
8.7249 | -91.0911
8.7252 | -91.0974
8.7255 | -91.1003
8.7258 | -91.1029
8.7261 | -91.1092
8.7264 | -91.1154
8.7267 | -91.1197
8.727 | -91.1205
8.7273 | -91.1229
8.7276 | -91.1293
8.7279 | -91.1345
8.7282 | -91.1358
8.7286 | -91.1367
8.7289 | -91.1386
8.7292 | -91.1414
8.7295 | -91.1436
8.7298 | -91.1439
8.7301 | -91.146
8.7304 | -91.1472
8.7307 | -91.1486
8.731 | -91.1494
8.7313 | -91.1504
8.7316 | -91.1511
8.732 | -91.1517
8.7323 | -91.1524
8.7326 | -91.1535
8.7329 | -91.1541
8.7332 | -91.1546
8.7335 | -91.1556
8.7338 | -91.1574
8.7341 | -91.1584
8.7344 | -91.1592
8.7347 | -91.1601
8.735 | -91.1605
8.7353 | -91.1609
8.7357 | -91.1613
8.736 | -91.1624
8.7363 | -91.1633
8.7366 | -91.164
8.7369 | -91.1643
8.7372 | -91.1644
8.7375 | -91.1648
8.7378 | -91.1655
8.7381 | -91.166
8.7384 | -91.1664
8.7387 | -91.1668
8.7391 | -91.1675
8.7394 | -91.1688
8.7397 | -91.1696
8.74 | -91.17
8.7403 | -91.171
8.7406 | -91.1716
8.7409 | -91.1726
8.7412 | -91.1737
8.7415 | -91.174
8.7418 | -91.1746
8.7421 | -91.1754
8.7425 | -91.1764
8.7428 | -91.1772
8.7431 | -91.1783
8.7434 | -91.1788
8.7437 | -91.1801
8.744 | -91.1812
8.7443 | -91.1815
8.7446 | -91.1819
8.7449 | -91.1823
8.7452 | -91.1827
8.7455 | -91.1831
8.7458 | -91.1838
8.7462 | -91.1851
8.7465 | -91.1863
8.7468 | -91.1874
8.7471 | -91.1881
8.7474 | -91.1887
8.7477 | -91.1894
8.748 | -91.1901
8.7483 | -91.1914
8.7486 | -91.1919
8.7489 | -91.1925
8.7492 | -91.1932
8.7496 | -91.1936
8.7499 | -91.1943
8.7502 | -91.1949
8.7505 | -91.1955
8.7508 | -91.1958
8.7511 | -91.1964
8.7514 | -91.1966
8.7517 | -91.1963
8.752 | -91.1962
8.7523 | -91.1963
8.7526 | -91.197
8.7529 | -91.1978
8.7533 | -91.1984
8.7536 | -91.1992
8.7539 | -91.2002
8.7542 | -91.2008
8.7545 | -91.2016
8.7548 | -91.2023
8.7551 | -91.203
8.7554 | -91.2047
8.7557 | -91.206
8.756 | -91.2069
8.7563 | -91.2077
8.7567 | -91.2086
8.757 | -91.2098
8.7573 | -91.2103
8.7576 | -91.2108
8.7579 | -91.2113
8.7582 | -91.2125
8.7585 | -91.214
8.7588 | -91.2148
8.7591 | -91.2157
8.7594 | -91.2164
8.7597 | -91.2172
8.76 | -91.2179
8.7604 | -91.2187
8.7607 | -91.2195
8.761 | -91.2205
8.7613 | -91.222
8.7616 | -91.2231
8.7619 | -91.2239
8.7622 | -91.2246
8.7625 | -91.2256
8.7628 | -91.2266
8.7631 | -91.227
8.7634 | -91.2276
8.7638 | -91.2279
8.7641 | -91.2287
8.7644 | -91.2294
8.7647 | -91.2301
8.765 | -91.2308
8.7653 | -91.2324
8.7656 | -91.2335
8.7659 | -91.2341
8.7662 | -91.2345
8.7665 | -91.2353
8.7668 | -91.2359
8.7672 | -91.2366
8.7675 | -91.2375
8.7678 | -91.2389
8.7681 | -91.2409
8.7684 | -91.2421
8.7687 | -91.243
8.769 | -91.2445
8.7693 | -91.2455
8.7696 | -91.2461
8.7699 | -91.2472
8.7702 | -91.2485
8.7705 | -91.2494
8.7709 | -91.2499
8.7712 | -91.2503
8.7715 | -91.2506
8.7718 | -91.2513
8.7721 | -91.2523
8.7724 | -91.2532
8.7727 | -91.2543
8.773 | -91.2551
8.7733 | -91.2557
8.7736 | -91.2562
8.7739 | -91.2571
8.7743 | -91.2583
8.7746 | -91.2595
8.7749 | -91.2609
8.7752 | -91.2617
8.7755 | -91.2626
8.7758 | -91.2634
8.7761 | -91.2642
8.7764 | -91.2652
8.7767 | -91.2665
8.777 | -91.2676
8.7773 | -91.2684
8.7776 | -91.2695
8.778 | -91.2709
8.7783 | -91.2722
8.7786 | -91.2732
8.7789 | -91.2736
8.7792 | -91.2744
8.7795 | -91.275
8.7798 | -91.2758
8.7801 | -91.2766
8.7804 | -91.2776
8.7807 | -91.2787
8.781 | -91.2794
8.7814 | -91.2799
8.7817 | -91.2808
8.782 | -91.2829
8.7823 | -91.2846
8.7826 | -91.2853
8.7829 | -91.2858
8.7832 | -91.2865
8.7835 | -91.2871
8.7838 | -91.2883
8.7841 | -91.2891
8.7844 | -91.2905
8.7848 | -91.2916
8.7851 | -91.2923
8.7854 | -91.293
8.7857 | -91.2948
8.786 | -91.2972
8.7863 | -91.2988
8.7866 | -91.2995
8.7869 | -91.3003
8.7872 | -91.3008
8.7875 | -91.3019
8.7878 | -91.3035
8.7881 | -91.3042
8.7885 | -91.305
8.7888 | -91.3058
8.7891 | -91.3062
8.7894 | -91.3071
8.7897 | -91.3081
8.79 | -91.3099
8.7903 | -91.3124
8.7906 | -91.3133
8.7909 | -91.3141
8.7912 | -91.3152
8.7915 | -91.3166
8.7919 | -91.319
8.7922 | -91.321
8.7925 | -91.3218
8.7928 | -91.3229
8.7931 | -91.324
8.7934 | -91.325
8.7937 | -91.3261
8.794 | -91.3275
8.7943 | -91.3283
8.7946 | -91.3296
8.7949 | -91.3304
8.7952 | -91.3315
8.7956 | -91.3327
8.7959 | -91.3339
8.7962 | -91.3355
8.7965 | -91.337
8.7968 | -91.3383
8.7971 | -91.3391
8.7974 | -91.3403
8.7977 | -91.3414
8.798 | -91.3421
8.7983 | -91.3427
8.7986 | -91.3439
8.799 | -91.3453
8.7993 | -91.3462
8.7996 | -91.3471
8.7999 | -91.3477
8.8002 | -91.3483
8.8005 | -91.349
8.8008 | -91.3498
8.8011 | -91.3507
8.8014 | -91.3514
8.8017 | -91.3527
8.802 | -91.3535
8.8024 | -91.3541
8.8027 | -91.3549
8.803 | -91.3559
8.8033 | -91.3567
8.8036 | -91.3579
8.8039 | -91.3587
8.8042 | -91.3598
8.8045 | -91.3604
8.8048 | -91.3615
8.8051 | -91.3626
8.8054 | -91.3637
8.8057 | -91.3653
8.8061 | -91.3667
8.8064 | -91.3683
8.8067 | -91.3703
8.807 | -91.3716
8.8073 | -91.3724
8.8076 | -91.3733
8.8079 | -91.3739
8.8082 | -91.3749
8.8085 | -91.3764
8.8088 | -91.3775
8.8091 | -91.3782
8.8095 | -91.3787
8.8098 | -91.3798
8.8101 | -91.3812
8.8104 | -91.3818
8.8107 | -91.3828
8.811 | -91.3836
8.8113 | -91.3852
8.8116 | -91.3861
8.8119 | -91.3872
8.8122 | -91.3883
8.8125 | -91.3892
8.8128 | -91.3903
8.8132 | -91.392
8.8135 | -91.3932
8.8138 | -91.394
8.8141 | -91.3951
8.8144 | -91.396
8.8147 | -91.3969
8.815 | -91.398
8.8153 | -91.3988
8.8156 | -91.4
8.8159 | -91.4015
8.8162 | -91.4024
8.8166 | -91.4037
8.8169 | -91.4051
8.8172 | -91.406
8.8175 | -91.4068
8.8178 | -91.4075
8.8181 | -91.4078
8.8184 | -91.4081
8.8187 | -91.4088
8.819 | -91.4097
8.8193 | -91.411
8.8196 | -91.4115
8.82 | -91.4126
8.8203 | -91.4134
8.8206 | -91.414
8.8209 | -91.4147
8.8212 | -91.4158
8.8215 | -91.4171
8.8218 | -91.4188
8.8221 | -91.4206
8.8224 | -91.4213
8.8227 | -91.4216
8.823 | -91.4225
8.8233 | -91.4239
8.8237 | -91.4248
8.824 | -91.4253
8.8243 | -91.4259
8.8246 | -91.4267
8.8249 | -91.4278
8.8252 | -91.429
8.8255 | -91.4301
8.8258 | -91.4308
8.8261 | -91.4312
8.8264 | -91.4328
8.8267 | -91.4339
8.8271 | -91.4348
8.8274 | -91.4359
8.8277 | -91.4366
8.828 | -91.4375
8.8283 | -91.4388
8.8286 | -91.4399
8.8289 | -91.4411
8.8292 | -91.4424
8.8295 | -91.4434
8.8298 | -91.4445
8.8301 | -91.4458
8.8304 | -91.4468
8.8308 | -91.4477
8.8311 | -91.4487
8.8314 | -91.4495
8.8317 | -91.4501
8.832 | -91.4514
8.8323 | -91.4526
8.8326 | -91.4538
8.8329 | -91.4553
8.8332 | -91.4564
8.8335 | -91.4576
8.8338 | -91.4601
8.8342 | -91.4617
8.8345 | -91.4626
8.8348 | -91.4632
8.8351 | -91.4641
8.8354 | -91.465
8.8357 | -91.466
8.836 | -91.4672
8.8363 | -91.4682
8.8366 | -91.4693
8.8369 | -91.4701
8.8372 | -91.4711
8.8376 | -91.4719
8.8379 | -91.4732
8.8382 | -91.4743
8.8385 | -91.4755
8.8388 | -91.476
8.8391 | -91.4775
8.8394 | -91.4794
8.8397 | -91.4811
8.84 | -91.4821
8.8403 | -91.4831
8.8406 | -91.4843
8.8409 | -91.4852
8.8413 | -91.4866
8.8416 | -91.4881
8.8419 | -91.4891
8.8422 | -91.4899
8.8425 | -91.4906
8.8428 | -91.4919
8.8431 | -91.4929
8.8434 | -91.4938
8.8437 | -91.4947
8.844 | -91.4954
8.8443 | -91.4963
8.8447 | -91.4973
8.845 | -91.4981
8.8453 | -91.4988
8.8456 | -91.4997
8.8459 | -91.5008
8.8462 | -91.5023
8.8465 | -91.5046
8.8468 | -91.5057
8.8471 | -91.5068
8.8474 | -91.5075
8.8477 | -91.5081
8.848 | -91.5092
8.8484 | -91.51
8.8487 | -91.5112
8.849 | -91.5125
8.8493 | -91.5138
8.8496 | -91.5149
8.8499 | -91.5158
8.8502 | -91.5168
8.8505 | -91.5177
8.8508 | -91.5186
8.8511 | -91.5198
8.8514 | -91.5215
8.8518 | -91.5228
8.8521 | -91.524
8.8524 | -91.525
8.8527 | -91.5261
8.853 | -91.528
8.8533 | -91.5295
8.8536 | -91.5303
8.8539 | -91.5311
8.8542 | -91.5318
8.8545 | -91.5324
8.8548 | -91.533
8.8552 | -91.5331
8.8555 | -91.5343
8.8558 | -91.5352
8.8561 | -91.5363
8.8564 | -91.5373
8.8567 | -91.5382
8.857 | -91.5392
8.8573 | -91.5403
8.8576 | -91.5415
8.8579 | -91.5428
8.8582 | -91.5438
8.8585 | -91.5446
8.8589 | -91.5453
8.8592 | -91.5458
8.8595 | -91.5468
8.8598 | -91.5477
8.8601 | -91.5487
8.8604 | -91.5499
8.8607 | -91.5514
8.861 | -91.5526
8.8613 | -91.5536
8.8616 | -91.5548
8.8619 | -91.5559
8.8623 | -91.5569
8.8626 | -91.5579
8.8629 | -91.5588
8.8632 | -91.56
8.8635 | -91.5616
8.8638 | -91.5624
8.8641 | -91.5636
8.8644 | -91.5646
8.8647 | -91.5653
8.865 | -91.566
8.8653 | -91.5669
8.8656 | -91.5675
8.866 | -91.5683
8.8663 | -91.5694
8.8666 | -91.5706
8.8669 | -91.5713
8.8672 | -91.5724
8.8675 | -91.5738
8.8678 | -91.575
8.8681 | -91.5764
8.8684 | -91.5773
8.8687 | -91.5782
8.869 | -91.5794
8.8694 | -91.5805
8.8697 | -91.5813
8.87 | -91.582
8.8703 | -91.5832
8.8706 | -91.5844
8.8709 | -91.5857
8.8712 | -91.5867
8.8715 | -91.5878
8.8718 | -91.589
8.8721 | -91.5904
8.8724 | -91.5918
8.8728 | -91.5931
8.8731 | -91.5946
8.8734 | -91.5957
8.8737 | -91.5969
8.874 | -91.598
8.8743 | -91.5991
8.8746 | -91.6002
8.8749 | -91.6013
8.8752 | -91.6027
8.8755 | -91.6041
8.8758 | -91.6053
8.8761 | -91.6062
8.8765 | -91.6074
8.8768 | -91.6085
8.8771 | -91.6098
8.8774 | -91.6118
8.8777 | -91.6128
8.878 | -91.6139
8.8783 | -91.6152
8.8786 | -91.6165
8.8789 | -91.6179
8.8792 | -91.619
8.8795 | -91.6196
8.8799 | -91.6199
8.8802 | -91.6211
8.8805 | -91.6237
8.8808 | -91.6263
8.8811 | -91.6275
8.8814 | -91.6285
8.8817 | -91.6299
8.882 | -91.6308
8.8823 | -91.6319
8.8826 | -91.6329
8.8829 | -91.634
8.8832 | -91.635
8.8836 | -91.6361
8.8839 | -91.6368
8.8842 | -91.6377
8.8845 | -91.6387
8.8848 | -91.6399
8.8851 | -91.6408
8.8854 | -91.6414
8.8857 | -91.6423
8.886 | -91.6435
8.8863 | -91.6445
8.8866 | -91.6456
8.887 | -91.6469
8.8873 | -91.6479
8.8876 | -91.649
8.8879 | -91.6503
8.8882 | -91.6513
8.8885 | -91.6523
8.8888 | -91.6531
8.8891 | -91.655
8.8894 | -91.6571
8.8897 | -91.6585
8.89 | -91.6594
8.8903 | -91.6604
8.8907 | -91.6619
8.891 | -91.6627
8.8913 | -91.6642
8.8916 | -91.6654
8.8919 | -91.6664
8.8922 | -91.6675
8.8925 | -91.6693
8.8928 | -91.6706
8.8931 | -91.6714
8.8934 | -91.6726
8.8937 | -91.6737
8.8941 | -91.6747
8.8944 | -91.676
8.8947 | -91.6774
8.895 | -91.6786
8.8953 | -91.6796
8.8956 | -91.6808
8.8959 | -91.6818
8.8962 | -91.6828
8.8965 | -91.6839
8.8968 | -91.6854
8.8971 | -91.6866
8.8975 | -91.6875
8.8978 | -91.6887
8.8981 | -91.6899
8.8984 | -91.6911
8.8987 | -91.6921
8.899 | -91.6929
8.8993 | -91.6937
8.8996 | -91.695
8.8999 | -91.6963
8.9002 | -91.6976
8.9005 | -91.6986
8.9008 | -91.6996
8.9012 | -91.7012
8.9015 | -91.7025
8.9018 | -91.7037
8.9021 | -91.7048
8.9024 | -91.706
8.9027 | -91.7076
8.903 | -91.7088
8.9033 | -91.7107
8.9036 | -91.7132
8.9039 | -91.7145
8.9042 | -91.7157
8.9046 | -91.7168
8.9049 | -91.718
8.9052 | -91.7192
8.9055 | -91.7203
8.9058 | -91.7216
8.9061 | -91.7234
8.9064 | -91.725
8.9067 | -91.7262
8.907 | -91.7272
8.9073 | -91.7284
8.9076 | -91.7294
8.9079 | -91.7299
8.9083 | -91.7308
8.9086 | -91.7323
8.9089 | -91.7336
8.9092 | -91.7348
8.9095 | -91.736
8.9098 | -91.7373
8.9101 | -91.7385
8.9104 | -91.7397
8.9107 | -91.7405
8.911 | -91.742
8.9113 | -91.7435
8.9117 | -91.7449
8.912 | -91.746
8.9123 | -91.7472
8.9126 | -91.7484
8.9129 | -91.7493
8.9132 | -91.7504
8.9135 | -91.7515
8.9138 | -91.7532
8.9141 | -91.7548
8.9144 | -91.7559
8.9147 | -91.7568
8.9151 | -91.7577
8.9154 | -91.7595
8.9157 | -91.7609
8.916 | -91.7625
8.9163 | -91.764
8.9166 | -91.7652
8.9169 | -91.7663
8.9172 | -91.7677
8.9175 | -91.7687
8.9178 | -91.7698
8.9181 | -91.7709
8.9184 | -91.7719
8.9188 | -91.7732
8.9191 | -91.7745
8.9194 | -91.7756
8.9197 | -91.7768
8.92 | -91.7771
8.9203 | -91.7778
8.9206 | -91.7789
8.9209 | -91.7799
8.9212 | -91.7808
8.9215 | -91.7825
8.9218 | -91.784
8.9222 | -91.7851
8.9225 | -91.7861
8.9228 | -91.7872
8.9231 | -91.7883
8.9234 | -91.7898
8.9237 | -91.7916
8.924 | -91.7933
8.9243 | -91.7944
8.9246 | -91.7956
8.9249 | -91.7965
8.9252 | -91.7977
8.9255 | -91.7988
8.9259 | -91.8
8.9262 | -91.8014
8.9265 | -91.803
8.9268 | -91.8043
8.9271 | -91.8055
8.9274 | -91.8066
8.9277 | -91.8074
8.928 | -91.8086
8.9283 | -91.8099
8.9286 | -91.8111
8.9289 | -91.8122
8.9293 | -91.8134
8.9296 | -91.8148
8.9299 | -91.8166
8.9302 | -91.8179
8.9305 | -91.8191
8.9308 | -91.8204
8.9311 | -91.8214
8.9314 | -91.8224
8.9317 | -91.8239
8.932 | -91.8258
8.9323 | -91.8269
8.9327 | -91.8278
8.933 | -91.829
8.9333 | -91.8302
8.9336 | -91.8314
8.9339 | -91.8326
8.9342 | -91.8337
8.9345 | -91.8346
8.9348 | -91.8359
8.9351 | -91.8372
8.9354 | -91.8387
8.9357 | -91.8397
8.936 | -91.8402
8.9364 | -91.8412
8.9367 | -91.842
8.937 | -91.8432
8.9373 | -91.8442
8.9376 | -91.8454
8.9379 | -91.8468
8.9382 | -91.8489
8.9385 | -91.8506
8.9388 | -91.8522
8.9391 | -91.8531
8.9394 | -91.8543
8.9398 | -91.8555
8.9401 | -91.8566
8.9404 | -91.8579
8.9407 | -91.8592
8.941 | -91.8608
8.9413 | -91.862
8.9416 | -91.863
8.9419 | -91.8645
8.9422 | -91.8658
8.9425 | -91.8672
8.9428 | -91.8689
8.9431 | -91.87
8.9435 | -91.8713
8.9438 | -91.8723
8.9441 | -91.8738
8.9444 | -91.8748
8.9447 | -91.8759
8.945 | -91.8769
8.9453 | -91.8779
8.9456 | -91.8791
8.9459 | -91.8811
8.9462 | -91.8819
8.9465 | -91.8828
8.9469 | -91.8841
8.9472 | -91.8854
8.9475 | -91.8868
8.9478 | -91.8879
8.9481 | -91.8891
8.9484 | -91.8906
8.9487 | -91.8925
8.949 | -91.8941
8.9493 | -91.8957
8.9496 | -91.8969
8.9499 | -91.8988
8.9503 | -91.9004
8.9506 | -91.902
8.9509 | -91.9032
8.9512 | -91.9043
8.9515 | -91.9053
8.9518 | -91.9065
8.9521 | -91.9075
8.9524 | -91.9091
8.9527 | -91.9107
8.953 | -91.9119
8.9533 | -91.913
8.9536 | -91.9141
8.954 | -91.9151
8.9543 | -91.9161
8.9546 | -91.917
8.9549 | -91.9179
8.9552 | -91.9192
8.9555 | -91.9204
8.9558 | -91.9216
8.9561 | -91.9227
8.9564 | -91.9239
8.9567 | -91.9251
8.957 | -91.9276
8.9574 | -91.9301
8.9577 | -91.9322
8.958 | -91.9334
8.9583 | -91.9348
8.9586 | -91.9361
8.9589 | -91.9373
8.9592 | -91.9386
8.9595 | -91.9399
8.9598 | -91.9412
8.9601 | -91.9425
8.9604 | -91.9439
8.9607 | -91.945
8.9611 | -91.9461
8.9614 | -91.9476
8.9617 | -91.9503
8.962 | -91.9512
8.9623 | -91.9522
8.9626 | -91.9533
8.9629 | -91.9547
8.9632 | -91.9561
8.9635 | -91.9581
8.9638 | -91.9597
8.9641 | -91.9608
8.9645 | -91.9623
8.9648 | -91.9634
8.9651 | -91.9646
8.9654 | -91.9659
8.9657 | -91.967
8.966 | -91.9684
8.9663 | -91.9696
8.9666 | -91.9708
8.9669 | -91.9721
8.9672 | -91.9732
8.9675 | -91.9743
8.9679 | -91.9751
8.9682 | -91.9768
8.9685 | -91.9784
8.9688 | -91.98
8.9691 | -91.9812
8.9694 | -91.9828
8.9697 | -91.9839
8.97 | -91.9851
8.9703 | -91.9864
8.9706 | -91.988
8.9709 | -91.9893
8.9712 | -91.9908
8.9716 | -91.992
8.9719 | -91.9932
8.9722 | -91.9947
8.9725 | -91.997
8.9728 | -92.0005
8.9731 | -92.0029
8.9734 | -92.0042
8.9737 | -92.0059
8.974 | -92.007
8.9743 | -92.0083
8.9746 | -92.0096
8.975 | -92.0102
8.9753 | -92.0128
8.9756 | -92.0157
8.9759 | -92.017
8.9762 | -92.0186
8.9765 | -92.0199
8.9768 | -92.0212
8.9771 | -92.0228
8.9774 | -92.0243
8.9777 | -92.0253
8.978 | -92.0265
8.9783 | -92.0279
8.9787 | -92.0295
8.979 | -92.0309
8.9793 | -92.032
8.9796 | -92.0336
8.9799 | -92.0349
8.9802 | -92.0363
8.9805 | -92.0378
8.9808 | -92.0392
8.9811 | -92.0406
8.9814 | -92.0419
8.9817 | -92.043
8.9821 | -92.0439
8.9824 | -92.0452
8.9827 | -92.0466
8.983 | -92.048
8.9833 | -92.0493
8.9836 | -92.0506
8.9839 | -92.0518
8.9842 | -92.0537
8.9845 | -92.0553
8.9848 | -92.0569
8.9851 | -92.0584
8.9855 | -92.0593
8.9858 | -92.0604
8.9861 | -92.0615
8.9864 | -92.0628
8.9867 | -92.0643
8.987 | -92.0647
8.9873 | -92.0647
8.9876 | -92.0656
8.9879 | -92.0666
8.9882 | -92.0678
8.9885 | -92.0691
8.9888 | -92.0706
8.9892 | -92.0726
8.9895 | -92.0751
8.9898 | -92.0763
8.9901 | -92.0775
8.9904 | -92.0785
8.9907 | -92.0795
8.991 | -92.0809
8.9913 | -92.0825
8.9916 | -92.0842
8.9919 | -92.0852
8.9922 | -92.0864
8.9926 | -92.0878
8.9929 | -92.0889
8.9932 | -92.0902
8.9935 | -92.0912
8.9938 | -92.0927
8.9941 | -92.094
8.9944 | -92.0949
8.9947 | -92.096
8.995 | -92.0969
8.9953 | -92.0977
8.9956 | -92.099
8.9959 | -92.1
8.9963 | -92.1013
8.9966 | -92.1026
8.9969 | -92.1037
8.9972 | -92.1049
8.9975 | -92.1057
8.9978 | -92.1072
8.9981 | -92.109
8.9984 | -92.1102
8.9987 | -92.1114
8.999 | -92.1125
8.9993 | -92.1138
8.9997 | -92.1145
9 | -92.1156
9.0003 | -92.1171
9.0006 | -92.1187
9.0009 | -92.12
9.0012 | -92.1214
9.0015 | -92.1223
9.0018 | -92.1228
9.0021 | -92.1238
9.0024 | -92.125
9.0027 | -92.1256
9.0031 | -92.1269
9.0034 | -92.128
9.0037 | -92.129
9.004 | -92.1301
9.0043 | -92.1315
9.0046 | -92.1332
9.0049 | -92.1347
9.0052 | -92.136
9.0055 | -92.1375
9.0058 | -92.1387
9.0061 | -92.1405
9.0064 | -92.142
9.0068 | -92.1437
9.0071 | -92.1452
9.0074 | -92.1469
9.0077 | -92.1489
9.008 | -92.1509
9.0083 | -92.1523
9.0086 | -92.1535
9.0089 | -92.1547
9.0092 | -92.1559
9.0095 | -92.1571
9.0098 | -92.1584
9.0102 | -92.1597
9.0105 | -92.1613
9.0108 | -92.1627
9.0111 | -92.164
9.0114 | -92.1655
9.0117 | -92.1675
9.012 | -92.1691
9.0123 | -92.1702
9.0126 | -92.1714
9.0129 | -92.1732
9.0132 | -92.1748
9.0135 | -92.1761
9.0139 | -92.1774
9.0142 | -92.1787
9.0145 | -92.1799
9.0148 | -92.1812
9.0151 | -92.1823
9.0154 | -92.1829
9.0157 | -92.1838
9.016 | -92.1849
9.0163 | -92.1858
9.0166 | -92.1871
9.0169 | -92.189
9.0173 | -92.1903
9.0176 | -92.1915
9.0179 | -92.1927
9.0182 | -92.1941
9.0185 | -92.1953
9.0188 | -92.1967
9.0191 | -92.1981
9.0194 | -92.1994
9.0197 | -92.2006
9.02 | -92.2018
9.0203 | -92.203
9.0206 | -92.205
9.021 | -92.2074
9.0213 | -92.2089
9.0216 | -92.21
9.0219 | -92.2113
9.0222 | -92.2124
9.0225 | -92.214
9.0228 | -92.2154
9.0231 | -92.2174
9.0234 | -92.2209
9.0237 | -92.2237
9.024 | -92.2248
9.0244 | -92.2259
9.0247 | -92.2272
9.025 | -92.2283
9.0253 | -92.2295
9.0256 | -92.2308
9.0259 | -92.2325
9.0262 | -92.2338
9.0265 | -92.2356
9.0268 | -92.2372
9.0271 | -92.2385
9.0274 | -92.2403
9.0278 | -92.2416
9.0281 | -92.244
9.0284 | -92.2461
9.0287 | -92.2475
9.029 | -92.2486
9.0293 | -92.2498
9.0296 | -92.2511
9.0299 | -92.2524
9.0302 | -92.2539
9.0305 | -92.2553
9.0308 | -92.2562
9.0311 | -92.2573
9.0315 | -92.2587
9.0318 | -92.26
9.0321 | -92.2612
9.0324 | -92.2623
9.0327 | -92.2633
9.033 | -92.2654
9.0333 | -92.2671
9.0336 | -92.2716
9.0339 | -92.2749
9.0342 | -92.2762
9.0345 | -92.2775
9.0349 | -92.2786
9.0352 | -92.2794
9.0355 | -92.2809
9.0358 | -92.283
9.0361 | -92.2857
9.0364 | -92.2882
9.0367 | -92.2902
9.037 | -92.2915
9.0373 | -92.2932
9.0376 | -92.2943
9.0379 | -92.2952
9.0382 | -92.2963
9.0386 | -92.2976
9.0389 | -92.2996
9.0392 | -92.3012
9.0395 | -92.3029
9.0398 | -92.3039
9.0401 | -92.3056
9.0404 | -92.3071
9.0407 | -92.3088
9.041 | -92.3099
9.0413 | -92.3113
9.0416 | -92.3126
9.042 | -92.3141
9.0423 | -92.3156
9.0426 | -92.3176
9.0429 | -92.3195
9.0432 | -92.3208
9.0435 | -92.3228
9.0438 | -92.3242
9.0441 | -92.3254
9.0444 | -92.3274
9.0447 | -92.3293
9.045 | -92.3308
9.0454 | -92.3327
9.0457 | -92.334
9.046 | -92.3353
9.0463 | -92.3362
9.0466 | -92.3372
9.0469 | -92.3385
9.0472 | -92.3398
9.0475 | -92.3411
9.0478 | -92.3429
9.0481 | -92.3444
9.0484 | -92.3455
9.0487 | -92.3468
9.0491 | -92.3481
9.0494 | -92.3495
9.0497 | -92.3508
9.05 | -92.3522
9.0503 | -92.3534
9.0506 | -92.3546
9.0509 | -92.356
9.0512 | -92.3572
9.0515 | -92.3593
9.0518 | -92.3618
9.0521 | -92.3633
9.0525 | -92.3645
9.0528 | -92.366
9.0531 | -92.3677
9.0534 | -92.369
9.0537 | -92.3699
9.054 | -92.3711
9.0543 | -92.3721
9.0546 | -92.3737
9.0549 | -92.3753
9.0552 | -92.3765
9.0555 | -92.3776
9.0558 | -92.3787
9.0562 | -92.3798
9.0565 | -92.381
9.0568 | -92.3823
9.0571 | -92.3834
9.0574 | -92.3845
9.0577 | -92.3852
9.058 | -92.3859
9.0583 | -92.3871
9.0586 | -92.388
9.0589 | -92.3892
9.0592 | -92.3906
9.0596 | -92.3919
9.0599 | -92.3936
9.0602 | -92.3947
9.0605 | -92.3957
9.0608 | -92.3967
9.0611 | -92.3983
9.0614 | -92.4005
9.0617 | -92.4024
9.062 | -92.4035
9.0623 | -92.4047
9.0626 | -92.4059
9.063 | -92.4071
9.0633 | -92.4086
9.0636 | -92.4101
9.0639 | -92.411
9.0642 | -92.4122
9.0645 | -92.4133
9.0648 | -92.4145
9.0651 | -92.4156
9.0654 | -92.4164
9.0657 | -92.4179
9.066 | -92.4195
9.0663 | -92.4203
9.0667 | -92.4213
9.067 | -92.4227
9.0673 | -92.4238
9.0676 | -92.4253
9.0679 | -92.4272
9.0682 | -92.4285
9.0685 | -92.43
9.0688 | -92.4317
9.0691 | -92.4329
9.0694 | -92.4351
9.0697 | -92.4366
9.0701 | -92.4377
9.0704 | -92.4387
9.0707 | -92.4402
9.071 | -92.4414
9.0713 | -92.4426
9.0716 | -92.4435
9.0719 | -92.4443
9.0722 | -92.4455
9.0725 | -92.4469
9.0728 | -92.4481
9.0731 | -92.4494
9.0734 | -92.4508
9.0738 | -92.4522
9.0741 | -92.4534
9.0744 | -92.4545
9.0747 | -92.4556
9.075 | -92.4568
9.0753 | -92.4583
9.0756 | -92.4592
9.0759 | -92.4604
9.0762 | -92.4614
9.0765 | -92.4617
9.0768 | -92.4625
9.0772 | -92.4638
9.0775 | -92.465
9.0778 | -92.4661
9.0781 | -92.4679
9.0784 | -92.4697
9.0787 | -92.4706
9.079 | -92.4716
9.0793 | -92.473
9.0796 | -92.4742
9.0799 | -92.4754
9.0802 | -92.4767
9.0806 | -92.4782
9.0809 | -92.4796
9.0812 | -92.481
9.0815 | -92.4824
9.0818 | -92.4834
9.0821 | -92.4845
9.0824 | -92.4858
9.0827 | -92.4872
9.083 | -92.4885
9.0833 | -92.4905
9.0836 | -92.4918
9.0839 | -92.4929
9.0843 | -92.4944
9.0846 | -92.4958
9.0849 | -92.4968
9.0852 | -92.498
9.0855 | -92.4992
9.0858 | -92.5009
9.0861 | -92.5029
9.0864 | -92.5045
9.0867 | -92.5059
9.087 | -92.5076
9.0873 | -92.5088
9.0877 | -92.5108
9.088 | -92.5134
9.0883 | -92.5145
9.0886 | -92.5154
9.0889 | -92.5163
9.0892 | -92.5174
9.0895 | -92.5186
9.0898 | -92.5194
9.0901 | -92.5201
9.0904 | -92.5211
9.0907 | -92.5223
9.091 | -92.5238
9.0914 | -92.525
9.0917 | -92.5261
9.092 | -92.5272
9.0923 | -92.5281
9.0926 | -92.529
9.0929 | -92.5304
9.0932 | -92.5317
9.0935 | -92.5333
9.0938 | -92.5347
9.0941 | -92.5358
9.0944 | -92.5368
9.0948 | -92.5379
9.0951 | -92.539
9.0954 | -92.5403
9.0957 | -92.5419
9.096 | -92.5431
9.0963 | -92.5439
9.0966 | -92.5449
9.0969 | -92.5461
9.0972 | -92.5472
9.0975 | -92.5484
9.0978 | -92.5495
9.0982 | -92.5507
9.0985 | -92.5519
9.0988 | -92.5531
9.0991 | -92.5546
9.0994 | -92.5558
9.0997 | -92.5569
9.1 | -92.5581
9.1003 | -92.5593
9.1006 | -92.5603
9.1009 | -92.5614
9.1012 | -92.5631
9.1015 | -92.564
9.1019 | -92.5651
9.1022 | -92.5667
9.1025 | -92.5692
9.1028 | -92.5705
9.1031 | -92.5714
9.1034 | -92.5726
9.1037 | -92.5737
9.104 | -92.575
9.1043 | -92.5759
9.1046 | -92.5766
9.1049 | -92.5779
9.1053 | -92.5806
9.1056 | -92.5828
9.1059 | -92.5839
9.1062 | -92.5855
9.1065 | -92.587
9.1068 | -92.5881
9.1071 | -92.5894
9.1074 | -92.5917
9.1077 | -92.5932
9.108 | -92.5947
9.1083 | -92.5964
9.1086 | -92.5976
9.109 | -92.5987
9.1093 | -92.6005
9.1096 | -92.6018
9.1099 | -92.6027
9.1102 | -92.6038
9.1105 | -92.6042
9.1108 | -92.6049
9.1111 | -92.6058
9.1114 | -92.6067
9.1117 | -92.6079
9.112 | -92.6091
9.1124 | -92.61
9.1127 | -92.611
9.113 | -92.6122
9.1133 | -92.6135
9.1136 | -92.6148
9.1139 | -92.6156
9.1142 | -92.6167
9.1145 | -92.6179
9.1148 | -92.6191
9.1151 | -92.6203
9.1154 | -92.6217
9.1158 | -92.6227
9.1161 | -92.6239
9.1164 | -92.625
9.1167 | -92.6262
9.117 | -92.6276
9.1173 | -92.6295
9.1176 | -92.631
9.1179 | -92.6322
9.1182 | -92.6333
9.1185 | -92.6347
9.1188 | -92.6358
9.1191 | -92.6371
9.1195 | -92.6388
9.1198 | -92.6404
9.1201 | -92.6416
9.1204 | -92.6429
9.1207 | -92.6442
9.121 | -92.6454
9.1213 | -92.6467
9.1216 | -92.6479
9.1219 | -92.6492
9.1222 | -92.6494
9.1225 | -92.6504
9.1229 | -92.6517
9.1232 | -92.6527
9.1235 | -92.6533
9.1238 | -92.6545
9.1241 | -92.6557
9.1244 | -92.657
9.1247 | -92.6582
9.125 | -92.6591
9.1253 | -92.6598
9.1256 | -92.6611
9.1259 | -92.6622
9.1262 | -92.6632
9.1266 | -92.6628
9.1269 | -92.6631
9.1272 | -92.6642
9.1275 | -92.6657
9.1278 | -92.6674
9.1281 | -92.6688
9.1284 | -92.67
9.1287 | -92.6713
9.129 | -92.6724
9.1293 | -92.674
9.1296 | -92.6753
9.13 | -92.6763
9.1303 | -92.6776
9.1306 | -92.6794
9.1309 | -92.6808
9.1312 | -92.6821
9.1315 | -92.6835
9.1318 | -92.6847
9.1321 | -92.6859
9.1324 | -92.6871
9.1327 | -92.6882
9.133 | -92.6888
9.1334 | -92.6898
9.1337 | -92.6911
9.134 | -92.6923
9.1343 | -92.6927
9.1346 | -92.6938
9.1349 | -92.6948
9.1352 | -92.6965
9.1355 | -92.6979
9.1358 | -92.6989
9.1361 | -92.7001
9.1364 | -92.7015
9.1367 | -92.7029
9.1371 | -92.7041
9.1374 | -92.7055
9.1377 | -92.7069
9.138 | -92.7084
9.1383 | -92.7094
9.1386 | -92.7103
9.1389 | -92.7111
9.1392 | -92.7122
9.1395 | -92.7133
9.1398 | -92.7139
9.1401 | -92.7138
9.1405 | -92.7158
9.1408 | -92.7178
9.1411 | -92.7191
9.1414 | -92.7201
9.1417 | -92.7215
9.142 | -92.7228
9.1423 | -92.7238
9.1426 | -92.7251
9.1429 | -92.7256
9.1432 | -92.7272
9.1435 | -92.7285
9.1438 | -92.7292
9.1442 | -92.73
9.1445 | -92.7319
9.1448 | -92.7353
9.1451 | -92.737
9.1454 | -92.7384
9.1457 | -92.7393
9.146 | -92.74
9.1463 | -92.741
9.1466 | -92.7419
9.1469 | -92.7428
9.1472 | -92.7433
9.1476 | -92.7441
9.1479 | -92.7451
9.1482 | -92.746
9.1485 | -92.7474
9.1488 | -92.7493
9.1491 | -92.7504
9.1494 | -92.752
9.1497 | -92.753
9.15 | -92.754
9.1503 | -92.7545
9.1506 | -92.7555
9.1509 | -92.7566
9.1513 | -92.7577
9.1516 | -92.7587
9.1519 | -92.7599
9.1522 | -92.7614
9.1525 | -92.763
9.1528 | -92.7645
9.1531 | -92.7657
9.1534 | -92.7674
9.1537 | -92.7691
9.154 | -92.7708
9.1543 | -92.7722
9.1547 | -92.7733
9.155 | -92.7746
9.1553 | -92.7757
9.1556 | -92.7767
9.1559 | -92.7779
9.1562 | -92.7791
9.1565 | -92.7796
9.1568 | -92.7807
9.1571 | -92.7817
9.1574 | -92.7826
9.1577 | -92.7838
9.1581 | -92.7849
9.1584 | -92.7861
9.1587 | -92.7874
9.159 | -92.7883
9.1593 | -92.7895
9.1596 | -92.7911
9.1599 | -92.7919
9.1602 | -92.7932
9.1605 | -92.7942
9.1608 | -92.7955
9.1611 | -92.7966
9.1614 | -92.7983
9.1618 | -92.7994
9.1621 | -92.7993
9.1624 | -92.8
9.1627 | -92.801
9.163 | -92.8019
9.1633 | -92.8025
9.1636 | -92.8037
9.1639 | -92.8052
9.1642 | -92.8072
9.1645 | -92.8086
9.1648 | -92.809
9.1652 | -92.8091
9.1655 | -92.8099
9.1658 | -92.811
9.1661 | -92.8119
9.1664 | -92.8135
9.1667 | -92.8146
9.167 | -92.8157
9.1673 | -92.8167
9.1676 | -92.8186
9.1679 | -92.8207
9.1682 | -92.8218
9.1685 | -92.8226
9.1689 | -92.8235
9.1692 | -92.8246
9.1695 | -92.8252
9.1698 | -92.8262
9.1701 | -92.8279
9.1704 | -92.8295
9.1707 | -92.8307
9.171 | -92.8324
9.1713 | -92.835
9.1716 | -92.8371
9.1719 | -92.8381
9.1723 | -92.8389
9.1726 | -92.8399
9.1729 | -92.841
9.1732 | -92.8423
9.1735 | -92.843
9.1738 | -92.8442
9.1741 | -92.8458
9.1744 | -92.8466
9.1747 | -92.8474
9.175 | -92.8486
9.1753 | -92.8501
9.1757 | -92.852
9.176 | -92.8531
9.1763 | -92.854
9.1766 | -92.8548
9.1769 | -92.8551
9.1772 | -92.8556
9.1775 | -92.856
9.1778 | -92.8561
9.1781 | -92.8563
9.1784 | -92.857
9.1787 | -92.8582
9.179 | -92.8597
9.1794 | -92.8611
9.1797 | -92.8626
9.18 | -92.8641
9.1803 | -92.8651
9.1806 | -92.8662
9.1809 | -92.8671
9.1812 | -92.8678
9.1815 | -92.8687
9.1818 | -92.8696
9.1821 | -92.871
9.1824 | -92.8728
9.1828 | -92.8741
9.1831 | -92.8752
9.1834 | -92.8759
9.1837 | -92.8764
9.184 | -92.8775
9.1843 | -92.8788
9.1846 | -92.8797
9.1849 | -92.8807
9.1852 | -92.8814
9.1855 | -92.8826
9.1858 | -92.884
9.1861 | -92.8854
9.1865 | -92.8869
9.1868 | -92.888
9.1871 | -92.8887
9.1874 | -92.8897
9.1877 | -92.8918
9.188 | -92.893
9.1883 | -92.8941
9.1886 | -92.8953
9.1889 | -92.8962
9.1892 | -92.897
9.1895 | -92.8982
9.1899 | -92.9
9.1902 | -92.9009
9.1905 | -92.9013
9.1908 | -92.9024
9.1911 | -92.9032
9.1914 | -92.9041
9.1917 | -92.9047
9.192 | -92.9057
9.1923 | -92.9065
9.1926 | -92.9075
9.1929 | -92.9085
9.1933 | -92.9093
9.1936 | -92.9101
9.1939 | -92.9107
9.1942 | -92.9115
9.1945 | -92.9127
9.1948 | -92.9136
9.1951 | -92.9143
9.1954 | -92.9155
9.1957 | -92.9163
9.196 | -92.9176
9.1963 | -92.9188
9.1966 | -92.919
9.197 | -92.9198
9.1973 | -92.9206
9.1976 | -92.9222
9.1979 | -92.9235
9.1982 | -92.9245
9.1985 | -92.9256
9.1988 | -92.926
9.1991 | -92.9261
9.1994 | -92.9265
9.1997 | -92.9272
9.2 | -92.9283
9.2004 | -92.9292
9.2007 | -92.93
9.201 | -92.931
9.2013 | -92.9323
9.2016 | -92.9334
9.2019 | -92.935
9.2022 | -92.9356
9.2025 | -92.9358
9.2028 | -92.9364
9.2031 | -92.9373
9.2034 | -92.9382
9.2037 | -92.939
9.2041 | -92.9397
9.2044 | -92.9407
9.2047 | -92.9417
9.205 | -92.9421
9.2053 | -92.9424
9.2056 | -92.943
9.2059 | -92.9433
9.2062 | -92.9441
9.2065 | -92.9449
9.2068 | -92.9457
9.2071 | -92.9469
9.2075 | -92.9475
9.2078 | -92.9477
9.2081 | -92.9484
9.2084 | -92.9493
9.2087 | -92.9509
9.209 | -92.9523
9.2093 | -92.953
9.2096 | -92.9539
9.2099 | -92.9554
9.2102 | -92.9562
9.2105 | -92.9568
9.2109 | -92.9581
9.2112 | -92.959
9.2115 | -92.9596
9.2118 | -92.9607
9.2121 | -92.9621
9.2124 | -92.9633
9.2127 | -92.9643
9.213 | -92.9649
9.2133 | -92.9652
9.2136 | -92.9658
9.2139 | -92.9664
9.2142 | -92.9665
9.2146 | -92.9673
9.2149 | -92.9683
9.2152 | -92.9692
9.2155 | -92.9699
9.2158 | -92.9706
9.2161 | -92.971
9.2164 | -92.9717
9.2167 | -92.9723
9.217 | -92.9732
9.2173 | -92.9732
9.2176 | -92.9739
9.218 | -92.9746
9.2183 | -92.975
9.2186 | -92.9755
9.2189 | -92.9769
9.2192 | -92.9779
9.2195 | -92.979
9.2198 | -92.9798
9.2201 | -92.98
9.2204 | -92.9803
9.2207 | -92.9811
9.221 | -92.9816
9.2213 | -92.9815
9.2217 | -92.982
9.222 | -92.9826
9.2223 | -92.9832
9.2226 | -92.984
9.2229 | -92.9848
9.2232 | -92.9853
9.2235 | -92.9854
9.2238 | -92.9862
9.2241 | -92.9874
9.2244 | -92.9883
9.2247 | -92.9891
9.2251 | -92.9901
9.2254 | -92.9906
9.2257 | -92.9913
9.226 | -92.992
9.2263 | -92.9929
9.2266 | -92.9934
9.2269 | -92.9941
9.2272 | -92.9944
9.2275 | -92.9946
9.2278 | -92.9955
9.2281 | -92.9961
9.2285 | -92.997
9.2288 | -92.9986
9.2291 | -93.0006
9.2294 | -93.0013
9.2297 | -93.002
9.23 | -93.0029
9.2303 | -93.0036
9.2306 | -93.0043
9.2309 | -93.0053
9.2312 | -93.006
9.2315 | -93.0063
9.2318 | -93.0069
9.2322 | -93.0075
9.2325 | -93.0078
9.2328 | -93.0075
9.2331 | -93.0079
9.2334 | -93.0084
9.2337 | -93.0091
9.234 | -93.0096
9.2343 | -93.0104
9.2346 | -93.0115
9.2349 | -93.0119
9.2352 | -93.0123
9.2356 | -93.0129
9.2359 | -93.0136
9.2362 | -93.0145
9.2365 | -93.0152
9.2368 | -93.0151
9.2371 | -93.0156
9.2374 | -93.0165
9.2377 | -93.0171
9.238 | -93.0178
9.2383 | -93.0181
9.2386 | -93.0187
9.2389 | -93.0193
9.2393 | -93.0195
9.2396 | -93.0203
9.2399 | -93.0209
9.2402 | -93.0214
9.2405 | -93.0221
9.2408 | -93.0228
9.2411 | -93.0238
9.2414 | -93.0238
9.2417 | -93.0234
9.242 | -93.0242
9.2423 | -93.0248
9.2427 | -93.0263
9.243 | -93.0275
9.2433 | -93.0282
9.2436 | -93.0287
9.2439 | -93.0289
9.2442 | -93.0288
9.2445 | -93.0288
9.2448 | -93.0296
9.2451 | -93.0297
9.2454 | -93.0296
9.2457 | -93.03
9.2461 | -93.0304
9.2464 | -93.0311
9.2467 | -93.0319
9.247 | -93.0325
9.2473 | -93.033
9.2476 | -93.0337
9.2479 | -93.0343
9.2482 | -93.0347
9.2485 | -93.0352
9.2488 | -93.0358
9.2491 | -93.0368
9.2494 | -93.0375
9.2498 | -93.0379
9.2501 | -93.039
9.2504 | -93.0407
9.2507 | -93.0418
9.251 | -93.0426
9.2513 | -93.0432
9.2516 | -93.0436
9.2519 | -93.044
9.2522 | -93.0446
9.2525 | -93.045
9.2528 | -93.0458
9.2532 | -93.0468
9.2535 | -93.0471
9.2538 | -93.0474
9.2541 | -93.048
9.2544 | -93.0486
9.2547 | -93.0492
9.255 | -93.0501
9.2553 | -93.0501
9.2556 | -93.05
9.2559 | -93.0497
9.2562 | -93.05
9.2565 | -93.0505
9.2569 | -93.0512
9.2572 | -93.0517
9.2575 | -93.0521
9.2578 | -93.0522
9.2581 | -93.0523
9.2584 | -93.0527
9.2587 | -93.0533
9.259 | -93.0541
9.2593 | -93.0544
9.2596 | -93.0546
9.2599 | -93.0547
9.2603 | -93.0548
9.2606 | -93.0552
9.2609 | -93.0553
9.2612 | -93.0562
9.2615 | -93.0568
9.2618 | -93.0574
9.2621 | -93.0579
9.2624 | -93.059
9.2627 | -93.0598
9.263 | -93.061
9.2633 | -93.0619
9.2637 | -93.0628
9.264 | -93.0636
9.2643 | -93.064
9.2646 | -93.0642
9.2649 | -93.0645
9.2652 | -93.0645
9.2655 | -93.0649
9.2658 | -93.0657
9.2661 | -93.0661
9.2664 | -93.0661
9.2667 | -93.0666
9.267 | -93.0665
9.2674 | -93.0668
9.2677 | -93.0674
9.268 | -93.0679
9.2683 | -93.0683
9.2686 | -93.069
9.2689 | -93.0693
9.2692 | -93.0691
9.2695 | -93.0693
9.2698 | -93.0697
9.2701 | -93.0699
9.2704 | -93.0703
9.2708 | -93.0707
9.2711 | -93.0705
9.2714 | -93.0708
9.2717 | -93.0714
9.272 | -93.0722
9.2723 | -93.0732
9.2726 | -93.0731
9.2729 | -93.0733
9.2732 | -93.0739
9.2735 | -93.0741
9.2738 | -93.0734
9.2741 | -93.0737
9.2745 | -93.0739
9.2748 | -93.0743
9.2751 | -93.0746
9.2754 | -93.0745
9.2757 | -93.0746
9.276 | -93.0757
9.2763 | -93.0757
9.2766 | -93.075
9.2769 | -93.0746
9.2772 | -93.0748
9.2775 | -93.0747
9.2779 | -93.0742
9.2782 | -93.0749
9.2785 | -93.0753
9.2788 | -93.0756
9.2791 | -93.076
9.2794 | -93.0766
9.2797 | -93.0778
9.28 | -93.08
9.2803 | -93.0801
9.2806 | -93.0802
9.2809 | -93.0806
9.2812 | -93.0799
9.2816 | -93.0767
9.2819 | -93.0762
9.2822 | -93.0764
9.2825 | -93.077
9.2828 | -93.0773
9.2831 | -93.0779
9.2834 | -93.078
9.2837 | -93.0775
9.284 | -93.0771
9.2843 | -93.0781
9.2846 | -93.0794
9.285 | -93.0802
9.2853 | -93.0805
9.2856 | -93.0801
9.2859 | -93.0803
9.2862 | -93.0806
9.2865 | -93.081
9.2868 | -93.0813
9.2871 | -93.081
9.2874 | -93.0807
9.2877 | -93.081
9.288 | -93.0811
9.2884 | -93.0811
9.2887 | -93.0812
9.289 | -93.082
9.2893 | -93.0824
9.2896 | -93.0826
9.2899 | -93.0824
9.2902 | -93.0828
9.2905 | -93.0826
9.2908 | -93.0827
9.2911 | -93.083
9.2914 | -93.0839
9.2917 | -93.0834
9.2921 | -93.085
9.2924 | -93.086
9.2927 | -93.0867
9.293 | -93.0871
9.2933 | -93.0877
9.2936 | -93.0886
9.2939 | -93.0897
9.2942 | -93.0901
9.2945 | -93.09
9.2948 | -93.09
9.2951 | -93.0893
9.2955 | -93.0881
9.2958 | -93.0877
9.2961 | -93.0876
9.2964 | -93.088
9.2967 | -93.0883
9.297 | -93.088
9.2973 | -93.0885
9.2976 | -93.0888
9.2979 | -93.0889
9.2982 | -93.0897
9.2985 | -93.0903
9.2988 | -93.0896
9.2992 | -93.089
9.2995 | -93.0889
9.2998 | -93.089
9.3001 | -93.0889
9.3004 | -93.0888
9.3007 | -93.0889
9.301 | -93.0892
9.3013 | -93.089
9.3016 | -93.0891
9.3019 | -93.0889
9.3022 | -93.0884
9.3026 | -93.0882
9.3029 | -93.0883
9.3032 | -93.0884
9.3035 | -93.0887
9.3038 | -93.0888
9.3041 | -93.0887
9.3044 | -93.0884
9.3047 | -93.0883
9.305 | -93.0884
9.3053 | -93.0879
9.3056 | -93.0875
9.306 | -93.0875
9.3063 | -93.088
9.3066 | -93.0872
9.3069 | -93.0867
9.3072 | -93.0864
9.3075 | -93.086
9.3078 | -93.0858
9.3081 | -93.0847
9.3084 | -93.0844
9.3087 | -93.0846
9.309 | -93.0845
9.3093 | -93.0847
9.3097 | -93.0849
9.31 | -93.0855
9.3103 | -93.0856
9.3106 | -93.0855
9.3109 | -93.0853
9.3112 | -93.0854
9.3115 | -93.0854
9.3118 | -93.085
9.3121 | -93.0849
9.3124 | -93.0849
9.3127 | -93.085
9.3131 | -93.0852
9.3134 | -93.0851
9.3137 | -93.0851
9.314 | -93.0856
9.3143 | -93.086
9.3146 | -93.0861
9.3149 | -93.0862
9.3152 | -93.0862
9.3155 | -93.0864
9.3158 | -93.0869
9.3161 | -93.0874
9.3164 | -93.0887
9.3168 | -93.0899
9.3171 | -93.0897
9.3174 | -93.0896
9.3177 | -93.0895
9.318 | -93.0893
9.3183 | -93.0891
9.3186 | -93.0892
9.3189 | -93.0892
9.3192 | -93.0891
9.3195 | -93.0892
9.3198 | -93.0891
9.3202 | -93.0893
9.3205 | -93.0898
9.3208 | -93.0898
9.3211 | -93.0898
9.3214 | -93.0898
9.3217 | -93.0901
9.322 | -93.0905
9.3223 | -93.0911
9.3226 | -93.0913
9.3229 | -93.0912
9.3232 | -93.0907
9.3236 | -93.0906
9.3239 | -93.0908
9.3242 | -93.0904
9.3245 | -93.0903
9.3248 | -93.0903
9.3251 | -93.0906
9.3254 | -93.091
9.3257 | -93.0914
9.326 | -93.0918
9.3263 | -93.0919
9.3266 | -93.0916
9.3269 | -93.0921
9.3273 | -93.0919
9.3276 | -93.0923
9.3279 | -93.0916
9.3282 | -93.0911
9.3285 | -93.0907
9.3288 | -93.0908
9.3291 | -93.0903
9.3294 | -93.0903
9.3297 | -93.0908
9.33 | -93.0907
9.3303 | -93.0902
9.3307 | -93.0905
9.331 | -93.0908
9.3313 | -93.0908
9.3316 | -93.0904
9.3319 | -93.0905
9.3322 | -93.0906
9.3325 | -93.0904
9.3328 | -93.0903
9.3331 | -93.0913
9.3334 | -93.0916
9.3337 | -93.0917
9.334 | -93.0921
9.3344 | -93.0933
9.3347 | -93.0943
9.335 | -93.0945
9.3353 | -93.094
9.3356 | -93.0938
9.3359 | -93.0936
9.3362 | -93.0932
9.3365 | -93.0937
9.3368 | -93.0936
9.3371 | -93.0935
9.3374 | -93.0933
9.3378 | -93.0928
9.3381 | -93.0927
9.3384 | -93.0922
9.3387 | -93.092
9.339 | -93.0914
9.3393 | -93.0912
9.3396 | -93.0911
9.3399 | -93.0908
9.3402 | -93.0906
9.3405 | -93.0905
9.3408 | -93.0906
9.3412 | -93.0901
9.3415 | -93.0892
9.3418 | -93.089
9.3421 | -93.0893
9.3424 | -93.0894
9.3427 | -93.0888
9.343 | -93.0881
9.3433 | -93.0879
9.3436 | -93.0881
9.3439 | -93.088
9.3442 | -93.0876
9.3445 | -93.0875
9.3449 | -93.0874
9.3452 | -93.0873
9.3455 | -93.087
9.3458 | -93.0869
9.3461 | -93.087
9.3464 | -93.0868
9.3467 | -93.0867
9.347 | -93.0864
9.3473 | -93.0855
9.3476 | -93.0847
9.3479 | -93.0839
9.3483 | -93.0833
9.3486 | -93.0831
9.3489 | -93.0828
9.3492 | -93.0822
9.3495 | -93.0813
9.3498 | -93.0803
9.3501 | -93.0793
9.3504 | -93.0788
9.3507 | -93.0786
9.351 | -93.0783
9.3513 | -93.078
9.3516 | -93.0778
9.352 | -93.0779
9.3523 | -93.0776
9.3526 | -93.0773
9.3529 | -93.0767
9.3532 | -93.0761
9.3535 | -93.0759
9.3538 | -93.0759
9.3541 | -93.075
9.3544 | -93.0754
9.3547 | -93.076
9.355 | -93.0755
9.3554 | -93.0754
9.3557 | -93.0755
9.356 | -93.0753
9.3563 | -93.0749
9.3566 | -93.074
9.3569 | -93.0742
9.3572 | -93.0744
9.3575 | -93.0744
9.3578 | -93.0743
9.3581 | -93.0731
9.3584 | -93.0727
9.3588 | -93.0724
9.3591 | -93.072
9.3594 | -93.072
9.3597 | -93.0716
9.36 | -93.0709
9.3603 | -93.0701
9.3606 | -93.0701
9.3609 | -93.0698
9.3612 | -93.0695
9.3615 | -93.0696
9.3618 | -93.0702
9.3621 | -93.0701
9.3625 | -93.0693
9.3628 | -93.0691
9.3631 | -93.0688
9.3634 | -93.0689
9.3637 | -93.069
9.364 | -93.0692
9.3643 | -93.0688
9.3646 | -93.068
9.3649 | -93.0678
9.3652 | -93.0682
9.3655 | -93.0681
9.3659 | -93.0672
9.3662 | -93.0668
9.3665 | -93.0664
9.3668 | -93.0658
9.3671 | -93.0659
9.3674 | -93.0656
9.3677 | -93.0649
9.368 | -93.0647
9.3683 | -93.0645
9.3686 | -93.0639
9.3689 | -93.0634
9.3692 | -93.063
9.3696 | -93.0626
9.3699 | -93.0627
9.3702 | -93.0623
9.3705 | -93.0624
9.3708 | -93.0623
9.3711 | -93.0623
9.3714 | -93.0624
9.3717 | -93.0616
9.372 | -93.0615
9.3723 | -93.0612
9.3726 | -93.0611
9.373 | -93.0616
9.3733 | -93.0616
9.3736 | -93.0618
9.3739 | -93.0618
9.3742 | -93.0614
9.3745 | -93.0608
9.3748 | -93.0603
9.3751 | -93.0596
9.3754 | -93.059
9.3757 | -93.0593
9.376 | -93.0589
9.3764 | -93.0583
9.3767 | -93.0578
9.377 | -93.0573
9.3773 | -93.0564
9.3776 | -93.0557
9.3779 | -93.0547
9.3782 | -93.0539
9.3785 | -93.0536
9.3788 | -93.0532
9.3791 | -93.0527
9.3794 | -93.0523
9.3797 | -93.052
9.3801 | -93.0516
9.3804 | -93.0525
9.3807 | -93.0522
9.381 | -93.0533
9.3813 | -93.054
9.3816 | -93.0545
9.3819 | -93.0543
9.3822 | -93.054
9.3825 | -93.0537
9.3828 | -93.0533
9.3831 | -93.0528
9.3835 | -93.0524
9.3838 | -93.0516
9.3841 | -93.0512
9.3844 | -93.0507
9.3847 | -93.0498
9.385 | -93.0493
9.3853 | -93.0485
9.3856 | -93.0476
9.3859 | -93.0471
9.3862 | -93.0466
9.3865 | -93.0464
9.3868 | -93.047
9.3872 | -93.0475
9.3875 | -93.0474
9.3878 | -93.0468
9.3881 | -93.0452
9.3884 | -93.0438
9.3887 | -93.0432
9.389 | -93.0427
9.3893 | -93.0416
9.3896 | -93.041
9.3899 | -93.0406
9.3902 | -93.04
9.3906 | -93.0396
9.3909 | -93.039
9.3912 | -93.0386
9.3915 | -93.0385
9.3918 | -93.0381
9.3921 | -93.0378
9.3924 | -93.0374
9.3927 | -93.0371
9.393 | -93.0368
9.3933 | -93.0367
9.3936 | -93.0368
9.394 | -93.0364
9.3943 | -93.0358
9.3946 | -93.0352
9.3949 | -93.0349
9.3952 | -93.0344
9.3955 | -93.0341
9.3958 | -93.0337
9.3961 | -93.0341
9.3964 | -93.0341
9.3967 | -93.0336
9.397 | -93.033
9.3973 | -93.0326
9.3977 | -93.0322
9.398 | -93.032
9.3983 | -93.032
9.3986 | -93.0316
9.3989 | -93.0311
9.3992 | -93.0308
9.3995 | -93.0302
9.3998 | -93.0296
9.4001 | -93.0294
9.4004 | -93.0288
9.4007 | -93.0281
9.4011 | -93.0273
9.4014 | -93.0271
9.4017 | -93.0271
9.402 | -93.0273
9.4023 | -93.027
9.4026 | -93.026
9.4029 | -93.0253
9.4032 | -93.0247
9.4035 | -93.024
9.4038 | -93.0237
9.4041 | -93.0227
9.4044 | -93.022
9.4048 | -93.0215
9.4051 | -93.021
9.4054 | -93.0203
9.4057 | -93.0193
9.406 | -93.0185
9.4063 | -93.018
9.4066 | -93.0171
9.4069 | -93.0162
9.4072 | -93.0148
9.4075 | -93.014
9.4078 | -93.0137
9.4082 | -93.0134
9.4085 | -93.0129
9.4088 | -93.0121
9.4091 | -93.0111
9.4094 | -93.0107
9.4097 | -93.0097
9.41 | -93.0091
9.4103 | -93.0089
9.4106 | -93.0082
9.4109 | -93.0074
9.4112 | -93.0068
9.4116 | -93.0064
9.4119 | -93.0058
9.4122 | -93.0055
9.4125 | -93.0054
9.4128 | -93.005
9.4131 | -93.0045
9.4134 | -93.0039
9.4137 | -93.0034
9.414 | -93.0034
9.4143 | -93.0031
9.4146 | -93.0029
9.4149 | -93.0024
9.4153 | -93.0017
9.4156 | -93.0009
9.4159 | -93
9.4162 | -92.9984
9.4165 | -92.998
9.4168 | -92.9977
9.4171 | -92.9971
9.4174 | -92.9964
9.4177 | -92.9959
9.418 | -92.9953
9.4183 | -92.995
9.4187 | -92.9949
9.419 | -92.9947
9.4193 | -92.9943
9.4196 | -92.9941
9.4199 | -92.994
9.4202 | -92.9933
9.4205 | -92.9927
9.4208 | -92.9924
9.4211 | -92.9919
9.4214 | -92.9912
9.4217 | -92.9907
9.422 | -92.9903
9.4224 | -92.9898
9.4227 | -92.9895
9.423 | -92.9891
9.4233 | -92.9884
9.4236 | -92.9879
9.4239 | -92.9875
9.4242 | -92.9867
9.4245 | -92.986
9.4248 | -92.9857
9.4251 | -92.9852
9.4254 | -92.9849
9.4258 | -92.9841
9.4261 | -92.9826
9.4264 | -92.9809
9.4267 | -92.9793
9.427 | -92.9785
9.4273 | -92.978
9.4276 | -92.9774
9.4279 | -92.977
9.4282 | -92.9766
9.4285 | -92.9762
9.4288 | -92.9758
9.4291 | -92.9748
9.4295 | -92.9736
9.4298 | -92.9723
9.4301 | -92.9713
9.4304 | -92.9714
9.4307 | -92.9708
9.431 | -92.9699
9.4313 | -92.9692
9.4316 | -92.9683
9.4319 | -92.9675
9.4322 | -92.9667
9.4325 | -92.9667
9.4329 | -92.9664
9.4332 | -92.9659
9.4335 | -92.9654
9.4338 | -92.9653
9.4341 | -92.9649
9.4344 | -92.9644
9.4347 | -92.9637
9.435 | -92.963
9.4353 | -92.9628
9.4356 | -92.9624
9.4359 | -92.9618
9.4363 | -92.9617
9.4366 | -92.9611
9.4369 | -92.9606
9.4372 | -92.9602
9.4375 | -92.9592
9.4378 | -92.9581
9.4381 | -92.9574
9.4384 | -92.9566
9.4387 | -92.9552
9.439 | -92.9549
9.4393 | -92.9545
9.4396 | -92.9538
9.44 | -92.9532
9.4403 | -92.9528
9.4406 | -92.9522
9.4409 | -92.9519
9.4412 | -92.9515
9.4415 | -92.951
9.4418 | -92.9504
9.4421 | -92.9501
9.4424 | -92.9503
9.4427 | -92.9499
9.443 | -92.9493
9.4434 | -92.949
9.4437 | -92.9493
9.444 | -92.9488
9.4443 | -92.9485
9.4446 | -92.9475
9.4449 | -92.9467
9.4452 | -92.9461
9.4455 | -92.9451
9.4458 | -92.9445
9.4461 | -92.944
9.4464 | -92.9434
9.4467 | -92.9429
9.4471 | -92.9421
9.4474 | -92.9416
9.4477 | -92.941
9.448 | -92.9401
9.4483 | -92.9387
9.4486 | -92.9377
9.4489 | -92.9368
9.4492 | -92.9362
9.4495 | -92.9355
9.4498 | -92.935
9.4501 | -92.9344
9.4505 | -92.9338
9.4508 | -92.9333
9.4511 | -92.9327
9.4514 | -92.9325
9.4517 | -92.9323
9.452 | -92.9318
9.4523 | -92.9319
9.4526 | -92.9315
9.4529 | -92.9309
9.4532 | -92.9304
9.4535 | -92.9302
9.4539 | -92.9299
9.4542 | -92.9292
9.4545 | -92.9287
9.4548 | -92.9284
9.4551 | -92.9281
9.4554 | -92.9275
9.4557 | -92.9267
9.456 | -92.9265
9.4563 | -92.9261
9.4566 | -92.926
9.4569 | -92.9262
9.4572 | -92.926
9.4576 | -92.9252
9.4579 | -92.9248
9.4582 | -92.9249
9.4585 | -92.9249
9.4588 | -92.9244
9.4591 | -92.9238
9.4594 | -92.9236
9.4597 | -92.923
9.46 | -92.9226
9.4603 | -92.9227
9.4606 | -92.9221
9.461 | -92.9216
9.4613 | -92.921
9.4616 | -92.9202
9.4619 | -92.92
9.4622 | -92.9197
9.4625 | -92.9184
9.4628 | -92.9166
9.4631 | -92.9152
9.4634 | -92.9141
9.4637 | -92.9131
9.464 | -92.9124
9.4643 | -92.9118
9.4647 | -92.911
9.465 | -92.9096
9.4653 | -92.9089
9.4656 | -92.9083
9.4659 | -92.9078
9.4662 | -92.9074
9.4665 | -92.9075
9.4668 | -92.9071
9.4671 | -92.9058
9.4674 | -92.9048
9.4677 | -92.9043
9.4681 | -92.9037
9.4684 | -92.9028
9.4687 | -92.9023
9.469 | -92.9019
9.4693 | -92.9011
9.4696 | -92.901
9.4699 | -92.9008
9.4702 | -92.9006
9.4705 | -92.9
9.4708 | -92.8994
9.4711 | -92.8987
9.4715 | -92.8979
9.4718 | -92.8976
9.4721 | -92.8975
9.4724 | -92.8966
9.4727 | -92.8961
9.473 | -92.8957
9.4733 | -92.8954
9.4736 | -92.8949
9.4739 | -92.894
9.4742 | -92.8934
9.4745 | -92.8926
9.4748 | -92.8918
9.4752 | -92.8912
9.4755 | -92.891
9.4758 | -92.8906
9.4761 | -92.89
9.4764 | -92.8895
9.4767 | -92.8889
9.477 | -92.8884
9.4773 | -92.8878
9.4776 | -92.8877
9.4779 | -92.8873
9.4782 | -92.8866
9.4786 | -92.886
9.4789 | -92.8854
9.4792 | -92.8846
9.4795 | -92.8838
9.4798 | -92.8831
9.4801 | -92.8825
9.4804 | -92.8821
9.4807 | -92.8816
9.481 | -92.8814
9.4813 | -92.8811
9.4816 | -92.8806
9.4819 | -92.88
9.4823 | -92.8793
9.4826 | -92.8787
9.4829 | -92.8783
9.4832 | -92.8774
9.4835 | -92.8769
9.4838 | -92.8773
9.4841 | -92.877
9.4844 | -92.8767
9.4847 | -92.8761
9.485 | -92.8755
9.4853 | -92.8745
9.4857 | -92.8737
9.486 | -92.8727
9.4863 | -92.8719
9.4866 | -92.8715
9.4869 | -92.8711
9.4872 | -92.8705
9.4875 | -92.8693
9.4878 | -92.8684
9.4881 | -92.8678
9.4884 | -92.867
9.4887 | -92.8663
9.4891 | -92.866
9.4894 | -92.8656
9.4897 | -92.8648
9.49 | -92.8642
9.4903 | -92.8639
9.4906 | -92.8634
9.4909 | -92.8632
9.4912 | -92.8631
9.4915 | -92.8637
9.4918 | -92.8635
9.4921 | -92.8629
9.4924 | -92.8622
9.4928 | -92.8612
9.4931 | -92.8607
9.4934 | -92.8602
9.4937 | -92.8599
9.494 | -92.8593
9.4943 | -92.8585
9.4946 | -92.8579
9.4949 | -92.857
9.4952 | -92.8565
9.4955 | -92.857
9.4958 | -92.8571
9.4962 | -92.8565
9.4965 | -92.856
9.4968 | -92.8553
9.4971 | -92.8548
9.4974 | -92.8542
9.4977 | -92.8536
9.498 | -92.8528
9.4983 | -92.852
9.4986 | -92.851
9.4989 | -92.8503
9.4992 | -92.8498
9.4995 | -92.8493
9.4999 | -92.8488
9.5002 | -92.8484
9.5005 | -92.8479
9.5008 | -92.8475
9.5011 | -92.847
9.5014 | -92.8461
9.5017 | -92.8455
9.502 | -92.8456
9.5023 | -92.8451
9.5026 | -92.8446
9.5029 | -92.8447
9.5033 | -92.8443
9.5036 | -92.8435
9.5039 | -92.8429
9.5042 | -92.8428
9.5045 | -92.8424
9.5048 | -92.842
9.5051 | -92.8415
9.5054 | -92.8407
9.5057 | -92.84
9.506 | -92.8397
9.5063 | -92.8393
9.5067 | -92.8389
9.507 | -92.8384
9.5073 | -92.8381
9.5076 | -92.8379
9.5079 | -92.8371
9.5082 | -92.8366
9.5085 | -92.8363
9.5088 | -92.8361
9.5091 | -92.8357
9.5094 | -92.8347
9.5097 | -92.8338
9.51 | -92.8332
9.5104 | -92.8325
9.5107 | -92.8319
9.511 | -92.8311
9.5113 | -92.8303
9.5116 | -92.8291
9.5119 | -92.8279
9.5122 | -92.8273
9.5125 | -92.8269
9.5128 | -92.8268
9.5131 | -92.8261
9.5134 | -92.8253
9.5138 | -92.8245
9.5141 | -92.8237
9.5144 | -92.8235
9.5147 | -92.8228
9.515 | -92.8223
9.5153 | -92.8217
9.5156 | -92.8213
9.5159 | -92.821
9.5162 | -92.8207
9.5165 | -92.8192
9.5168 | -92.8176
9.5171 | -92.8169
9.5175 | -92.8166
9.5178 | -92.8171
9.5181 | -92.8175
9.5184 | -92.8171
9.5187 | -92.8165
9.519 | -92.8161
9.5193 | -92.8159
9.5196 | -92.8155
9.5199 | -92.815
9.5202 | -92.8145
9.5205 | -92.8144
9.5209 | -92.8138
9.5212 | -92.8134
9.5215 | -92.813
9.5218 | -92.8127
9.5221 | -92.8123
9.5224 | -92.8121
9.5227 | -92.8117
9.523 | -92.8112
9.5233 | -92.8114
9.5236 | -92.811
9.5239 | -92.8108
9.5243 | -92.8107
9.5246 | -92.8102
9.5249 | -92.8097
9.5252 | -92.8094
9.5255 | -92.8094
9.5258 | -92.8089
9.5261 | -92.8082
9.5264 | -92.8077
9.5267 | -92.8072
9.527 | -92.8069
9.5273 | -92.8069
9.5276 | -92.8066
9.528 | -92.8064
9.5283 | -92.806
9.5286 | -92.8054
9.5289 | -92.8052
9.5292 | -92.8051
9.5295 | -92.8046
9.5298 | -92.8042
9.5301 | -92.8038
9.5304 | -92.8029
9.5307 | -92.8025
9.531 | -92.8022
9.5314 | -92.8018
9.5317 | -92.8013
9.532 | -92.8007
9.5323 | -92.7999
9.5326 | -92.7998
9.5329 | -92.7997
9.5332 | -92.7994
9.5335 | -92.7986
9.5338 | -92.7978
9.5341 | -92.7975
9.5344 | -92.7971
9.5347 | -92.7967
9.5351 | -92.7963
9.5354 | -92.7959
9.5357 | -92.7956
9.536 | -92.7954
9.5363 | -92.7947
9.5366 | -92.7941
9.5369 | -92.7938
9.5372 | -92.7932
9.5375 | -92.793
9.5378 | -92.7928
9.5381 | -92.7925
9.5385 | -92.7924
9.5388 | -92.7924
9.5391 | -92.792
9.5394 | -92.7915
9.5397 | -92.7909
9.54 | -92.7905
9.5403 | -92.7902
9.5406 | -92.7899
9.5409 | -92.79
9.5412 | -92.7895
9.5415 | -92.7892
9.5419 | -92.7887
9.5422 | -92.7883
9.5425 | -92.7878
9.5428 | -92.787
9.5431 | -92.7861
9.5434 | -92.7858
9.5437 | -92.785
9.544 | -92.7841
9.5443 | -92.7836
9.5446 | -92.7832
9.5449 | -92.7828
9.5452 | -92.7826
9.5456 | -92.7822
9.5459 | -92.7818
9.5462 | -92.7806
9.5465 | -92.7802
9.5468 | -92.7801
9.5471 | -92.7797
9.5474 | -92.7795
9.5477 | -92.7794
9.548 | -92.7797
9.5483 | -92.7794
9.5486 | -92.7789
9.549 | -92.7786
9.5493 | -92.7784
9.5496 | -92.7782
9.5499 | -92.7785
9.5502 | -92.7784
9.5505 | -92.7782
9.5508 | -92.7779
9.5511 | -92.7776
9.5514 | -92.7774
9.5517 | -92.7772
9.552 | -92.777
9.5523 | -92.7768
9.5527 | -92.7765
9.553 | -92.776
9.5533 | -92.7759
9.5536 | -92.7756
9.5539 | -92.7749
9.5542 | -92.7746
9.5545 | -92.7744
9.5548 | -92.774
9.5551 | -92.7735
9.5554 | -92.7736
9.5557 | -92.7739
9.5561 | -92.7732
9.5564 | -92.7723
9.5567 | -92.7719
9.557 | -92.7716
9.5573 | -92.7709
9.5576 | -92.7705
9.5579 | -92.7704
9.5582 | -92.7703
9.5585 | -92.7706
9.5588 | -92.7707
9.5591 | -92.7707
9.5595 | -92.7706
9.5598 | -92.7705
9.5601 | -92.7702
9.5604 | -92.7699
9.5607 | -92.7696
9.561 | -92.7692
9.5613 | -92.77
9.5616 | -92.7703
9.5619 | -92.7701
9.5622 | -92.7703
9.5625 | -92.7701
9.5628 | -92.7698
9.5632 | -92.7695
9.5635 | -92.7692
9.5638 | -92.7688
9.5641 | -92.7681
9.5644 | -92.7678
9.5647 | -92.7676
9.565 | -92.7674
9.5653 | -92.7672
9.5656 | -92.7669
9.5659 | -92.7668
9.5662 | -92.767
9.5666 | -92.767
9.5669 | -92.767
9.5672 | -92.7668
9.5675 | -92.7666
9.5678 | -92.7663
9.5681 | -92.7661
9.5684 | -92.7659
9.5687 | -92.7659
9.569 | -92.7657
9.5693 | -92.7656
9.5696 | -92.7658
9.5699 | -92.7659
9.5703 | -92.7658
9.5706 | -92.7656
9.5709 | -92.7653
9.5712 | -92.7651
9.5715 | -92.765
9.5718 | -92.7648
9.5721 | -92.7646
9.5724 | -92.7645
9.5727 | -92.7643
9.573 | -92.7645
9.5733 | -92.7644
9.5737 | -92.7642
9.574 | -92.7641
9.5743 | -92.7645
9.5746 | -92.7649
9.5749 | -92.7647
9.5752 | -92.7647
9.5755 | -92.7647
9.5758 | -92.7643
9.5761 | -92.7635
9.5764 | -92.7629
9.5767 | -92.7616
9.577 | -92.7602
9.5774 | -92.7591
9.5777 | -92.7588
9.578 | -92.7585
9.5783 | -92.7588
9.5786 | -92.7593
9.5789 | -92.7594
9.5792 | -92.7599
9.5795 | -92.76
9.5798 | -92.7599
9.5801 | -92.7597
9.5804 | -92.7597
9.5808 | -92.7596
9.5811 | -92.7597
9.5814 | -92.7595
9.5817 | -92.7597
9.582 | -92.7597
9.5823 | -92.7596
9.5826 | -92.7593
9.5829 | -92.7589
9.5832 | -92.7587
9.5835 | -92.7583
9.5838 | -92.7582
9.5842 | -92.7585
9.5845 | -92.7581
9.5848 | -92.758
9.5851 | -92.7586
9.5854 | -92.7587
9.5857 | -92.7592
9.586 | -92.7592
9.5863 | -92.7591
9.5866 | -92.7591
9.5869 | -92.759
9.5872 | -92.7587
9.5875 | -92.7581
9.5879 | -92.757
9.5882 | -92.7571
9.5885 | -92.7576
9.5888 | -92.7577
9.5891 | -92.7575
9.5894 | -92.758
9.5897 | -92.7584
9.59 | -92.7581
9.5903 | -92.758
9.5906 | -92.7579
9.5909 | -92.7578
9.5913 | -92.7579
9.5916 | -92.7578
9.5919 | -92.7574
9.5922 | -92.7577
9.5925 | -92.7577
9.5928 | -92.7576
9.5931 | -92.7576
9.5934 | -92.7573
9.5937 | -92.7562
9.594 | -92.7551
9.5943 | -92.7545
9.5946 | -92.755
9.595 | -92.7551
9.5953 | -92.755
9.5956 | -92.7553
9.5959 | -92.7554
9.5962 | -92.7554
9.5965 | -92.7554
9.5968 | -92.7557
9.5971 | -92.7557
9.5974 | -92.7555
9.5977 | -92.7557
9.598 | -92.7557
9.5984 | -92.7561
9.5987 | -92.7565
9.599 | -92.7565
9.5993 | -92.7561
9.5996 | -92.7559
9.5999 | -92.756
9.6002 | -92.7555
9.6005 | -92.7552
9.6008 | -92.7551
9.6011 | -92.7552
9.6014 | -92.7553
9.6018 | -92.7555
9.6021 | -92.7556
9.6024 | -92.7556
9.6027 | -92.7557
9.603 | -92.7556
9.6033 | -92.7559
9.6036 | -92.7556
9.6039 | -92.7556
9.6042 | -92.7557
9.6045 | -92.7558
9.6048 | -92.756
9.6051 | -92.7561
9.6055 | -92.7561
9.6058 | -92.7561
9.6061 | -92.7562
9.6064 | -92.757
9.6067 | -92.7587
9.607 | -92.7591
9.6073 | -92.7591
9.6076 | -92.7591
9.6079 | -92.7593
9.6082 | -92.7596
9.6085 | -92.7597
9.6089 | -92.7601
9.6092 | -92.7604
9.6095 | -92.7606
9.6098 | -92.7607
9.6101 | -92.7603
9.6104 | -92.7598
9.6107 | -92.7598
9.611 | -92.76
9.6113 | -92.76
9.6116 | -92.7599
9.6119 | -92.7602
9.6122 | -92.7611
9.6126 | -92.7615
9.6129 | -92.7619
9.6132 | -92.7625
9.6135 | -92.7625
9.6138 | -92.7621
9.6141 | -92.7618
9.6144 | -92.762
9.6147 | -92.7617
9.615 | -92.7617
9.6153 | -92.7624
9.6156 | -92.7626
9.616 | -92.7628
9.6163 | -92.7627
9.6166 | -92.7629
9.6169 | -92.7629
9.6172 | -92.7632
9.6175 | -92.7634
9.6178 | -92.7633
9.6181 | -92.7629
9.6184 | -92.763
9.6187 | -92.763
9.619 | -92.7629
9.6194 | -92.7631
9.6197 | -92.7634
9.62 | -92.7643
9.6203 | -92.765
9.6206 | -92.7652
9.6209 | -92.7655
9.6212 | -92.7665
9.6215 | -92.7666
9.6218 | -92.7667
9.6221 | -92.7668
9.6224 | -92.7669
9.6227 | -92.7672
9.6231 | -92.7676
9.6234 | -92.7676
9.6237 | -92.7677
9.624 | -92.7679
9.6243 | -92.7679
9.6246 | -92.7681
9.6249 | -92.7683
9.6252 | -92.7686
9.6255 | -92.7688
9.6258 | -92.7692
9.6261 | -92.7697
9.6265 | -92.7706
9.6268 | -92.7708
9.6271 | -92.7707
9.6274 | -92.7711
9.6277 | -92.7717
9.628 | -92.772
9.6283 | -92.7722
9.6286 | -92.7726
9.6289 | -92.773
9.6292 | -92.7733
9.6295 | -92.7735
9.6298 | -92.7737
9.6302 | -92.7746
9.6305 | -92.7748
9.6308 | -92.775
9.6311 | -92.7753
9.6314 | -92.7759
9.6317 | -92.7763
9.632 | -92.7767
9.6323 | -92.7769
9.6326 | -92.7769
9.6329 | -92.778
9.6332 | -92.7794
9.6336 | -92.7799
9.6339 | -92.7802
9.6342 | -92.7803
9.6345 | -92.7804
9.6348 | -92.7808
9.6351 | -92.7812
9.6354 | -92.7814
9.6357 | -92.7817
9.636 | -92.7823
9.6363 | -92.7828
9.6366 | -92.7834
9.637 | -92.7841
9.6373 | -92.7844
9.6376 | -92.7848
9.6379 | -92.7851
9.6382 | -92.7857
9.6385 | -92.7861
9.6388 | -92.7867
9.6391 | -92.7867
9.6394 | -92.7869
9.6397 | -92.7872
9.64 | -92.7878
9.6403 | -92.7883
9.6407 | -92.7888
9.641 | -92.7893
9.6413 | -92.7896
9.6416 | -92.7903
9.6419 | -92.7913
9.6422 | -92.7921
9.6425 | -92.793
9.6428 | -92.7937
9.6431 | -92.7939
9.6434 | -92.7944
9.6437 | -92.7948
9.6441 | -92.7954
9.6444 | -92.7962
9.6447 | -92.7967
9.645 | -92.797
9.6453 | -92.7973
9.6456 | -92.7971
9.6459 | -92.7972
9.6462 | -92.7975
9.6465 | -92.7983
9.6468 | -92.7993
9.6471 | -92.7998
9.6474 | -92.7998
9.6478 | -92.8001
9.6481 | -92.8013
9.6484 | -92.8017
9.6487 | -92.8026
9.649 | -92.8039
9.6493 | -92.8046
9.6496 | -92.8051
9.6499 | -92.8066
9.6502 | -92.8078
9.6505 | -92.8085
9.6508 | -92.809
9.6512 | -92.8095
9.6515 | -92.81
9.6518 | -92.8103
9.6521 | -92.8104
9.6524 | -92.8106
9.6527 | -92.8109
9.653 | -92.8112
9.6533 | -92.8115
9.6536 | -92.812
9.6539 | -92.8122
9.6542 | -92.8118
9.6546 | -92.8118
9.6549 | -92.8121
9.6552 | -92.8125
9.6555 | -92.8129
9.6558 | -92.8132
9.6561 | -92.8131
9.6564 | -92.8127
9.6567 | -92.8129
9.657 | -92.8137
9.6573 | -92.8136
9.6576 | -92.8138
9.6579 | -92.8144
9.6583 | -92.8146
9.6586 | -92.815
9.6589 | -92.8155
9.6592 | -92.8158
9.6595 | -92.8164
9.6598 | -92.8173
9.6601 | -92.8184
9.6604 | -92.8189
9.6607 | -92.8192
9.661 | -92.8193
9.6613 | -92.8199
9.6617 | -92.8204
9.662 | -92.8208
9.6623 | -92.821
9.6626 | -92.8215
9.6629 | -92.822
9.6632 | -92.8231
9.6635 | -92.8237
9.6638 | -92.8246
9.6641 | -92.8251
9.6644 | -92.8255
9.6647 | -92.826
9.665 | -92.8261
9.6654 | -92.8263
9.6657 | -92.8276
9.666 | -92.8279
9.6663 | -92.8283
9.6666 | -92.8293
9.6669 | -92.8303
9.6672 | -92.8308
9.6675 | -92.8313
9.6678 | -92.8315
9.6681 | -92.8315
9.6684 | -92.8317
9.6688 | -92.8323
9.6691 | -92.8327
9.6694 | -92.8331
9.6697 | -92.8335
9.67 | -92.8341
9.6703 | -92.8349
9.6706 | -92.8355
9.6709 | -92.8358
9.6712 | -92.8362
9.6715 | -92.8366
9.6718 | -92.8372
9.6722 | -92.8374
9.6725 | -92.8384
9.6728 | -92.8385
9.6731 | -92.8386
9.6734 | -92.8391
9.6737 | -92.8397
9.674 | -92.8403
9.6743 | -92.8404
9.6746 | -92.8402
9.6749 | -92.8405
9.6752 | -92.841
9.6755 | -92.8415
9.6759 | -92.842
9.6762 | -92.8424
9.6765 | -92.8428
9.6768 | -92.8433
9.6771 | -92.8437
9.6774 | -92.844
9.6777 | -92.8446
9.678 | -92.8446
9.6783 | -92.845
9.6786 | -92.8456
9.6789 | -92.8465
9.6793 | -92.8469
9.6796 | -92.8475
9.6799 | -92.8479
9.6802 | -92.8479
9.6805 | -92.8483
9.6808 | -92.8487
9.6811 | -92.8492
9.6814 | -92.85
9.6817 | -92.8509
9.682 | -92.8516
9.6823 | -92.8521
9.6826 | -92.8527
9.683 | -92.8531
9.6833 | -92.8537
9.6836 | -92.8547
9.6839 | -92.8552
9.6842 | -92.8554
9.6845 | -92.8561
9.6848 | -92.8569
9.6851 | -92.8575
9.6854 | -92.8577
9.6857 | -92.8581
9.686 | -92.8586
9.6864 | -92.8592
9.6867 | -92.8598
9.687 | -92.8606
9.6873 | -92.8615
9.6876 | -92.8626
9.6879 | -92.8636
9.6882 | -92.8645
9.6885 | -92.8652
9.6888 | -92.8661
9.6891 | -92.867
9.6894 | -92.8676
9.6898 | -92.8682
9.6901 | -92.8691
9.6904 | -92.8698
9.6907 | -92.8703
9.691 | -92.8707
9.6913 | -92.8714
9.6916 | -92.8723
9.6919 | -92.873
9.6922 | -92.8736
9.6925 | -92.8746
9.6928 | -92.8754
9.6931 | -92.8761
9.6935 | -92.8774
9.6938 | -92.8781
9.6941 | -92.8791
9.6944 | -92.8799
9.6947 | -92.8802
9.695 | -92.8808
9.6953 | -92.8815
9.6956 | -92.8821
9.6959 | -92.8829
9.6962 | -92.8838
9.6965 | -92.8848
9.6969 | -92.8856
9.6972 | -92.8862
9.6975 | -92.8869
9.6978 | -92.8875
9.6981 | -92.8882
9.6984 | -92.8891
9.6987 | -92.8902
9.699 | -92.8908
9.6993 | -92.8914
9.6996 | -92.8921
9.6999 | -92.8931
9.7002 | -92.8939
9.7006 | -92.8947
9.7009 | -92.8954
9.7012 | -92.896
9.7015 | -92.8962
9.7018 | -92.897
9.7021 | -92.898
9.7024 | -92.8988
9.7027 | -92.8995
9.703 | -92.8997
9.7033 | -92.8996
9.7036 | -92.8997
9.704 | -92.9003
9.7043 | -92.9011
9.7046 | -92.9021
9.7049 | -92.9028
9.7052 | -92.9039
9.7055 | -92.9047
9.7058 | -92.9054
9.7061 | -92.906
9.7064 | -92.9067
9.7067 | -92.9075
9.707 | -92.9087
9.7073 | -92.9091
9.7077 | -92.9096
9.708 | -92.9102
9.7083 | -92.9111
9.7086 | -92.912
9.7089 | -92.9125
9.7092 | -92.9132
9.7095 | -92.914
9.7098 | -92.9145
9.7101 | -92.9152
9.7104 | -92.9154
9.7107 | -92.916
9.7111 | -92.9168
9.7114 | -92.9174
9.7117 | -92.9182
9.712 | -92.9191
9.7123 | -92.9201
9.7126 | -92.9213
9.7129 | -92.9224
9.7132 | -92.9234
9.7135 | -92.9244
9.7138 | -92.9252
9.7141 | -92.926
9.7145 | -92.9267
9.7148 | -92.9275
9.7151 | -92.9282
9.7154 | -92.9294
9.7157 | -92.9302
9.716 | -92.9311
9.7163 | -92.9318
9.7166 | -92.9327
9.7169 | -92.934
9.7172 | -92.935
9.7175 | -92.9359
9.7178 | -92.9373
9.7182 | -92.9382
9.7185 | -92.9389
9.7188 | -92.94
9.7191 | -92.9408
9.7194 | -92.9406
9.7197 | -92.9413
9.72 | -92.9426
9.7203 | -92.9434
9.7206 | -92.9442
9.7209 | -92.945
9.7212 | -92.9458
9.7216 | -92.9466
9.7219 | -92.9475
9.7222 | -92.9484
9.7225 | -92.9494
9.7228 | -92.9502
9.7231 | -92.9513
9.7234 | -92.9523
9.7237 | -92.9536
9.724 | -92.9548
9.7243 | -92.9557
9.7246 | -92.9565
9.7249 | -92.957
9.7253 | -92.9571
9.7256 | -92.9574
9.7259 | -92.9576
9.7262 | -92.9581
9.7265 | -92.9587
9.7268 | -92.9592
9.7271 | -92.9594
9.7274 | -92.9595
9.7277 | -92.9605
9.728 | -92.9616
9.7283 | -92.9622
9.7287 | -92.9633
9.729 | -92.9651
9.7293 | -92.9656
9.7296 | -92.9661
9.7299 | -92.9675
9.7302 | -92.9686
9.7305 | -92.9698
9.7308 | -92.9708
9.7311 | -92.9712
9.7314 | -92.9714
9.7317 | -92.9725
9.7321 | -92.9739
9.7324 | -92.9744
9.7327 | -92.9746
9.733 | -92.9752
9.7333 | -92.976
9.7336 | -92.9765
9.7339 | -92.9775
9.7342 | -92.9787
9.7345 | -92.9795
9.7348 | -92.98
9.7351 | -92.9807
9.7354 | -92.9817
9.7358 | -92.9822
9.7361 | -92.9821
9.7364 | -92.9822
9.7367 | -92.983
9.737 | -92.9839
9.7373 | -92.985
9.7376 | -92.9859
9.7379 | -92.9867
9.7382 | -92.9874
9.7385 | -92.9881
9.7388 | -92.9886
9.7392 | -92.9892
9.7395 | -92.9901
9.7398 | -92.991
9.7401 | -92.9919
9.7404 | -92.9929
9.7407 | -92.9936
9.741 | -92.994
9.7413 | -92.9945
9.7416 | -92.9954
9.7419 | -92.9957
9.7422 | -92.9962
9.7425 | -92.997
9.7429 | -92.9977
9.7432 | -92.999
9.7435 | -93
9.7438 | -93.0006
9.7441 | -93.001
9.7444 | -93.002
9.7447 | -93.0033
9.745 | -93.0041
9.7453 | -93.0049
9.7456 | -93.006
9.7459 | -93.007
9.7463 | -93.0079
9.7466 | -93.0086
9.7469 | -93.0094
9.7472 | -93.0101
9.7475 | -93.0106
9.7478 | -93.0115
9.7481 | -93.0125
9.7484 | -93.0133
9.7487 | -93.014
9.749 | -93.0145
9.7493 | -93.0149
9.7497 | -93.0157
9.75 | -93.0165
9.7503 | -93.0177
9.7506 | -93.0178
9.7509 | -93.0184
9.7512 | -93.0193
9.7515 | -93.0201
9.7518 | -93.021
9.7521 | -93.0218
9.7524 | -93.0227
9.7527 | -93.0232
9.753 | -93.0239
9.7534 | -93.0248
9.7537 | -93.0258
9.754 | -93.0262
9.7543 | -93.0267
9.7546 | -93.0279
9.7549 | -93.028
9.7552 | -93.0285
9.7555 | -93.0292
9.7558 | -93.0304
9.7561 | -93.0312
9.7564 | -93.0321
9.7568 | -93.033
9.7571 | -93.0339
9.7574 | -93.0347
9.7577 | -93.0353
9.758 | -93.0363
9.7583 | -93.0372
9.7586 | -93.0382
9.7589 | -93.0394
9.7592 | -93.0402
9.7595 | -93.0412
9.7598 | -93.042
9.7601 | -93.0429
9.7605 | -93.0443
9.7608 | -93.0452
9.7611 | -93.0459
9.7614 | -93.0473
9.7617 | -93.0489
9.762 | -93.049
9.7623 | -93.0489
9.7626 | -93.0491
9.7629 | -93.0496
9.7632 | -93.0503
9.7635 | -93.0511
9.7639 | -93.052
9.7642 | -93.0529
9.7645 | -93.0537
9.7648 | -93.0547
9.7651 | -93.0561
9.7654 | -93.0583
9.7657 | -93.0596
9.766 | -93.0604
9.7663 | -93.0612
9.7666 | -93.0619
9.7669 | -93.0628
9.7673 | -93.0638
9.7676 | -93.0649
9.7679 | -93.0658
9.7682 | -93.0669
9.7685 | -93.0679
9.7688 | -93.0686
9.7691 | -93.0695
9.7694 | -93.0705
9.7697 | -93.0713
9.77 | -93.0718
9.7703 | -93.0725
9.7706 | -93.0736
9.771 | -93.0744
9.7713 | -93.0751
9.7716 | -93.0758
9.7719 | -93.0759
9.7722 | -93.0764
9.7725 | -93.0771
9.7728 | -93.0774
9.7731 | -93.0784
9.7734 | -93.0793
9.7737 | -93.0802
9.774 | -93.081
9.7744 | -93.0818
9.7747 | -93.0826
9.775 | -93.0833
9.7753 | -93.0838
9.7756 | -93.0846
9.7759 | -93.0853
9.7762 | -93.0859
9.7765 | -93.0863
9.7768 | -93.087
9.7771 | -93.0874
9.7774 | -93.088
9.7777 | -93.089
9.7781 | -93.0903
9.7784 | -93.0913
9.7787 | -93.0921
9.779 | -93.093
9.7793 | -93.0943
9.7796 | -93.0951
9.7799 | -93.0959
9.7802 | -93.0966
9.7805 | -93.0972
9.7808 | -93.0982
9.7811 | -93.0994
9.7815 | -93.0997
9.7818 | -93.1003
9.7821 | -93.1011
9.7824 | -93.102
9.7827 | -93.1043
9.783 | -93.106
9.7833 | -93.1069
9.7836 | -93.1078
9.7839 | -93.1085
9.7842 | -93.1093
9.7845 | -93.1101
9.7849 | -93.1103
9.7852 | -93.1102
9.7855 | -93.1104
9.7858 | -93.111
9.7861 | -93.1114
9.7864 | -93.1119
9.7867 | -93.1128
9.787 | -93.1135
9.7873 | -93.1143
9.7876 | -93.1152
9.7879 | -93.1158
9.7882 | -93.1166
9.7886 | -93.1175
9.7889 | -93.1184
9.7892 | -93.12
9.7895 | -93.1215
9.7898 | -93.1221
9.7901 | -93.1229
9.7904 | -93.1239
9.7907 | -93.1246
9.791 | -93.1259
9.7913 | -93.127
9.7916 | -93.128
9.792 | -93.1289
9.7923 | -93.1299
9.7926 | -93.1308
9.7929 | -93.1314
9.7932 | -93.1319
9.7935 | -93.1331
9.7938 | -93.1341
9.7941 | -93.1351
9.7944 | -93.1358
9.7947 | -93.1365
9.795 | -93.1373
9.7953 | -93.1382
9.7957 | -93.1393
9.796 | -93.14
9.7963 | -93.1407
9.7966 | -93.1424
9.7969 | -93.1444
9.7972 | -93.1452
9.7975 | -93.1458
9.7978 | -93.1466
9.7981 | -93.1474
9.7984 | -93.1481
9.7987 | -93.1488
9.7991 | -93.1496
9.7994 | -93.1507
9.7997 | -93.1513
9.8 | -93.1517
9.8003 | -93.1529
9.8006 | -93.1537
9.8009 | -93.1545
9.8012 | -93.1557
9.8015 | -93.1564
9.8018 | -93.1567
9.8021 | -93.1571
9.8025 | -93.1578
9.8028 | -93.1582
9.8031 | -93.1592
9.8034 | -93.1601
9.8037 | -93.1603
9.804 | -93.1607
9.8043 | -93.1614
9.8046 | -93.1627
9.8049 | -93.1634
9.8052 | -93.164
9.8055 | -93.1648
9.8058 | -93.1653
9.8062 | -93.1661
9.8065 | -93.1665
9.8068 | -93.1673
9.8071 | -93.168
9.8074 | -93.1688
9.8077 | -93.1694
9.808 | -93.1704
9.8083 | -93.1711
9.8086 | -93.1719
9.8089 | -93.1727
9.8092 | -93.1733
9.8096 | -93.1739
9.8099 | -93.1744
9.8102 | -93.1751
9.8105 | -93.1759
9.8108 | -93.1766
9.8111 | -93.1776
9.8114 | -93.1784
9.8117 | -93.1791
9.812 | -93.1801
9.8123 | -93.1809
9.8126 | -93.1817
9.8129 | -93.1824
9.8133 | -93.183
9.8136 | -93.1831
9.8139 | -93.1835
9.8142 | -93.1842
9.8145 | -93.1855
9.8148 | -93.1863
9.8151 | -93.187
9.8154 | -93.1878
9.8157 | -93.1885
9.816 | -93.189
9.8163 | -93.1897
9.8167 | -93.1905
9.817 | -93.1908
9.8173 | -93.1908
9.8176 | -93.1915
9.8179 | -93.1923
9.8182 | -93.193
9.8185 | -93.1937
9.8188 | -93.1942
9.8191 | -93.1947
9.8194 | -93.1953
9.8197 | -93.1958
9.8201 | -93.1963
9.8204 | -93.1974
9.8207 | -93.1993
9.821 | -93.2003
9.8213 | -93.2009
9.8216 | -93.2019
9.8219 | -93.2034
9.8222 | -93.2042
9.8225 | -93.205
9.8228 | -93.2057
9.8231 | -93.2068
9.8234 | -93.2086
9.8238 | -93.2104
9.8241 | -93.2113
9.8244 | -93.2122
9.8247 | -93.2129
9.825 | -93.2136
9.8253 | -93.2144
9.8256 | -93.2151
9.8259 | -93.2157
9.8262 | -93.2155
9.8265 | -93.2162
9.8268 | -93.2169
9.8272 | -93.2176
9.8275 | -93.2183
9.8278 | -93.2189
9.8281 | -93.2197
9.8284 | -93.2205
9.8287 | -93.2216
9.829 | -93.2227
9.8293 | -93.2235
9.8296 | -93.224
9.8299 | -93.2247
9.8302 | -93.2255
9.8305 | -93.2264
9.8309 | -93.2271
9.8312 | -93.228
9.8315 | -93.229
9.8318 | -93.2297
9.8321 | -93.2305
9.8324 | -93.2313
9.8327 | -93.2323
9.833 | -93.2326
9.8333 | -93.2329
9.8336 | -93.2338
9.8339 | -93.235
9.8343 | -93.2358
9.8346 | -93.2365
9.8349 | -93.2369
9.8352 | -93.2374
9.8355 | -93.2381
9.8358 | -93.2388
9.8361 | -93.2395
9.8364 | -93.2401
9.8367 | -93.2409
9.837 | -93.2415
9.8373 | -93.2425
9.8377 | -93.2437
9.838 | -93.2452
9.8383 | -93.2462
9.8386 | -93.2467
9.8389 | -93.2474
9.8392 | -93.2478
9.8395 | -93.2482
9.8398 | -93.2486
9.8401 | -93.2489
9.8404 | -93.2495
9.8407 | -93.2503
9.841 | -93.2511
9.8414 | -93.2518
9.8417 | -93.252
9.842 | -93.2526
9.8423 | -93.2531
9.8426 | -93.2536
9.8429 | -93.2542
9.8432 | -93.2552
9.8435 | -93.2563
9.8438 | -93.2568
9.8441 | -93.2571
9.8444 | -93.2577
9.8448 | -93.2581
9.8451 | -93.2583
9.8454 | -93.2587
9.8457 | -93.2596
9.846 | -93.2606
9.8463 | -93.2615
9.8466 | -93.2622
9.8469 | -93.2628
9.8472 | -93.2633
9.8475 | -93.2639
9.8478 | -93.2648
9.8481 | -93.2655
9.8485 | -93.2663
9.8488 | -93.2668
9.8491 | -93.2677
9.8494 | -93.2684
9.8497 | -93.2689
9.85 | -93.2696
9.8503 | -93.2703
9.8506 | -93.2711
9.8509 | -93.2722
9.8512 | -93.2727
9.8515 | -93.2727
9.8519 | -93.273
9.8522 | -93.274
9.8525 | -93.2743
9.8528 | -93.275
9.8531 | -93.2758
9.8534 | -93.2763
9.8537 | -93.2766
9.854 | -93.2771
9.8543 | -93.2779
9.8546 | -93.2787
9.8549 | -93.2793
9.8552 | -93.2797
9.8556 | -93.2802
9.8559 | -93.281
9.8562 | -93.2819
9.8565 | -93.2829
9.8568 | -93.2833
9.8571 | -93.2838
9.8574 | -93.2843
9.8577 | -93.2853
9.858 | -93.286
9.8583 | -93.2867
9.8586 | -93.2873
9.859 | -93.288
9.8593 | -93.2887
9.8596 | -93.2893
9.8599 | -93.2903
9.8602 | -93.2917
9.8605 | -93.2927
9.8608 | -93.2936
9.8611 | -93.2945
9.8614 | -93.2952
9.8617 | -93.2957
9.862 | -93.2964
9.8624 | -93.2973
9.8627 | -93.2977
9.863 | -93.2982
9.8633 | -93.2993
9.8636 | -93.3002
9.8639 | -93.3009
9.8642 | -93.3018
9.8645 | -93.3028
9.8648 | -93.3032
9.8651 | -93.3037
9.8654 | -93.3041
9.8657 | -93.3048
9.8661 | -93.3055
9.8664 | -93.3061
9.8667 | -93.3072
9.867 | -93.3084
9.8673 | -93.3094
9.8676 | -93.3105
9.8679 | -93.3105
9.8682 | -93.311
9.8685 | -93.3115
9.8688 | -93.3119
9.8691 | -93.3126
9.8695 | -93.3135
9.8698 | -93.314
9.8701 | -93.3144
9.8704 | -93.3149
9.8707 | -93.3155
9.871 | -93.3161
9.8713 | -93.3165
9.8716 | -93.3168
9.8719 | -93.3167
9.8722 | -93.3163
9.8725 | -93.3168
9.8728 | -93.3173
9.8732 | -93.3178
9.8735 | -93.3184
9.8738 | -93.3189
9.8741 | -93.3196
9.8744 | -93.3203
9.8747 | -93.3209
9.875 | -93.3214
9.8753 | -93.3219
9.8756 | -93.3221
9.8759 | -93.3227
9.8762 | -93.3236
9.8766 | -93.3245
9.8769 | -93.3251
9.8772 | -93.3255
9.8775 | -93.3259
9.8778 | -93.3261
9.8781 | -93.3258
9.8784 | -93.3263
9.8787 | -93.327
9.879 | -93.3277
9.8793 | -93.3288
9.8796 | -93.3292
9.88 | -93.3305
9.8803 | -93.3311
9.8806 | -93.3317
9.8809 | -93.3326
9.8812 | -93.3336
9.8815 | -93.3341
9.8818 | -93.3346
9.8821 | -93.3351
9.8824 | -93.3356
9.8827 | -93.3359
9.883 | -93.3361
9.8833 | -93.337
9.8837 | -93.3377
9.884 | -93.3387
9.8843 | -93.3398
9.8846 | -93.3404
9.8849 | -93.3408
9.8852 | -93.3417
9.8855 | -93.3421
9.8858 | -93.3425
9.8861 | -93.3432
9.8864 | -93.3444
9.8867 | -93.3451
9.8871 | -93.3457
9.8874 | -93.3462
9.8877 | -93.3469
9.888 | -93.3475
9.8883 | -93.348
9.8886 | -93.3485
9.8889 | -93.3484
9.8892 | -93.3484
9.8895 | -93.3487
9.8898 | -93.3493
9.8901 | -93.3497
9.8904 | -93.3506
9.8908 | -93.3514
9.8911 | -93.3518
9.8914 | -93.3524
9.8917 | -93.3528
9.892 | -93.3532
9.8923 | -93.3538
9.8926 | -93.3543
9.8929 | -93.3546
9.8932 | -93.3553
9.8935 | -93.356
9.8938 | -93.3565
9.8942 | -93.3569
9.8945 | -93.357
9.8948 | -93.3573
9.8951 | -93.3581
9.8954 | -93.3588
9.8957 | -93.3596
9.896 | -93.3602
9.8963 | -93.3608
9.8966 | -93.3615
9.8969 | -93.3622
9.8972 | -93.3628
9.8976 | -93.3632
9.8979 | -93.3638
9.8982 | -93.3644
9.8985 | -93.3651
9.8988 | -93.3658
9.8991 | -93.3665
9.8994 | -93.367
9.8997 | -93.3674
9.9 | -93.3676
9.9003 | -93.368
9.9006 | -93.3691
9.9009 | -93.3696
9.9013 | -93.3699
9.9016 | -93.3703
9.9019 | -93.3708
9.9022 | -93.3709
9.9025 | -93.3711
9.9028 | -93.3718
9.9031 | -93.3724
9.9034 | -93.3729
9.9037 | -93.373
9.904 | -93.3738
9.9043 | -93.3737
9.9047 | -93.374
9.905 | -93.3746
9.9053 | -93.3751
9.9056 | -93.3765
9.9059 | -93.3779
9.9062 | -93.3784
9.9065 | -93.3786
9.9068 | -93.3789
9.9071 | -93.3795
9.9074 | -93.38
9.9077 | -93.3806
9.908 | -93.3811
9.9084 | -93.3819
9.9087 | -93.3833
9.909 | -93.3843
9.9093 | -93.3847
9.9096 | -93.3856
9.9099 | -93.3867
9.9102 | -93.3874
9.9105 | -93.3881
9.9108 | -93.3885
9.9111 | -93.389
9.9114 | -93.3895
9.9118 | -93.3899
9.9121 | -93.39
9.9124 | -93.3905
9.9127 | -93.3909
9.913 | -93.3914
9.9133 | -93.3914
9.9136 | -93.3918
9.9139 | -93.3921
9.9142 | -93.3924
9.9145 | -93.3928
9.9148 | -93.3931
9.9152 | -93.3929
9.9155 | -93.3933
9.9158 | -93.3938
9.9161 | -93.3943
9.9164 | -93.3953
9.9167 | -93.3956
9.917 | -93.396
9.9173 | -93.3964
9.9176 | -93.3967
9.9179 | -93.3972
9.9182 | -93.3976
9.9185 | -93.3982
9.9189 | -93.3982
9.9192 | -93.3983
9.9195 | -93.3988
9.9198 | -93.399
9.9201 | -93.3993
9.9204 | -93.4005
9.9207 | -93.4014
9.921 | -93.4017
9.9213 | -93.4025
9.9216 | -93.4034
9.9219 | -93.4039
9.9223 | -93.4044
9.9226 | -93.4045
9.9229 | -93.4046
9.9232 | -93.4052
9.9235 | -93.4058
9.9238 | -93.4061
9.9241 | -93.4064
9.9244 | -93.4069
9.9247 | -93.4072
9.925 | -93.4076
9.9253 | -93.4081
9.9256 | -93.4083
9.926 | -93.4087
9.9263 | -93.4095
9.9266 | -93.4099
9.9269 | -93.4099
9.9272 | -93.4101
9.9275 | -93.4106
9.9278 | -93.411
9.9281 | -93.4113
9.9284 | -93.4117
9.9287 | -93.4124
9.929 | -93.4129
9.9294 | -93.4132
9.9297 | -93.4132
9.93 | -93.4135
9.9303 | -93.414
9.9306 | -93.4143
9.9309 | -93.4143
9.9312 | -93.4145
9.9315 | -93.4151
9.9318 | -93.4153
9.9321 | -93.4154
9.9324 | -93.4163
9.9328 | -93.4174
9.9331 | -93.4164
9.9334 | -93.4168
9.9337 | -93.4174
9.934 | -93.4179
9.9343 | -93.4182
9.9346 | -93.4186
9.9349 | -93.4186
9.9352 | -93.4191
9.9355 | -93.4201
9.9358 | -93.4206
9.9361 | -93.421
9.9365 | -93.4215
9.9368 | -93.4216
9.9371 | -93.4216
9.9374 | -93.4223
9.9377 | -93.4224
9.938 | -93.4224
9.9383 | -93.422
9.9386 | -93.4219
9.9389 | -93.4232
9.9392 | -93.4238
9.9395 | -93.4242
9.9399 | -93.4248
9.9402 | -93.4252
9.9405 | -93.4262
9.9408 | -93.4264
9.9411 | -93.4268
9.9414 | -93.427
9.9417 | -93.4271
9.942 | -93.4258
9.9423 | -93.4263
9.9426 | -93.4268
9.9429 | -93.4272
9.9432 | -93.4281
9.9436 | -93.4298
9.9439 | -93.4314
9.9442 | -93.4318
9.9445 | -93.4323
9.9448 | -93.4334
9.9451 | -93.4333
9.9454 | -93.4339
9.9457 | -93.4341
9.946 | -93.4342
9.9463 | -93.4345
9.9466 | -93.4348
9.947 | -93.4353
9.9473 | -93.4365
9.9476 | -93.437
9.9479 | -93.4371
9.9482 | -93.438
9.9485 | -93.4423
9.9488 | -93.4438
9.9491 | -93.4441
9.9494 | -93.4444
9.9497 | -93.4444
9.95 | -93.4446
9.9504 | -93.445
9.9507 | -93.4455
9.951 | -93.446
9.9513 | -93.4473
9.9516 | -93.4486
9.9519 | -93.449
9.9522 | -93.449
9.9525 | -93.4492
9.9528 | -93.4499
9.9531 | -93.4511
9.9534 | -93.4515
9.9537 | -93.4517
9.9541 | -93.452
9.9544 | -93.4526
9.9547 | -93.4535
9.955 | -93.4542
9.9553 | -93.4544
9.9556 | -93.4549
9.9559 | -93.4554
9.9562 | -93.4555
9.9565 | -93.4553
9.9568 | -93.4554
9.9571 | -93.4556
9.9575 | -93.4562
9.9578 | -93.4561
9.9581 | -93.4564
9.9584 | -93.4572
9.9587 | -93.4579
9.959 | -93.4584
9.9593 | -93.4589
9.9596 | -93.4595
9.9599 | -93.4597
9.9602 | -93.4599
9.9605 | -93.4602
9.9608 | -93.4605
9.9612 | -93.4605
9.9615 | -93.4604
9.9618 | -93.4607
9.9621 | -93.4604
9.9624 | -93.4606
9.9627 | -93.4609
9.963 | -93.4613
9.9633 | -93.4623
9.9636 | -93.4642
9.9639 | -93.4652
9.9642 | -93.4656
9.9646 | -93.4655
9.9649 | -93.4658
9.9652 | -93.4654
9.9655 | -93.4647
9.9658 | -93.4646
9.9661 | -93.4647
9.9664 | -93.4649
9.9667 | -93.4652
9.967 | -93.4662
9.9673 | -93.4668
9.9676 | -93.4669
9.968 | -93.4672
9.9683 | -93.4674
9.9686 | -93.4666
9.9689 | -93.4665
9.9692 | -93.467
9.9695 | -93.4674
9.9698 | -93.4678
9.9701 | -93.4684
9.9704 | -93.4682
9.9707 | -93.4682
9.971 | -93.469
9.9713 | -93.4691
9.9717 | -93.469
9.972 | -93.4692
9.9723 | -93.4694
9.9726 | -93.4696
9.9729 | -93.4701
9.9732 | -93.4703
9.9735 | -93.4704
9.9738 | -93.4706
9.9741 | -93.4716
9.9744 | -93.4726
9.9747 | -93.4732
9.9751 | -93.4735
9.9754 | -93.4747
9.9757 | -93.4759
9.976 | -93.4764
9.9763 | -93.4761
9.9766 | -93.4766
9.9769 | -93.4779
9.9772 | -93.4783
9.9775 | -93.4788
9.9778 | -93.4789
9.9781 | -93.4784
9.9784 | -93.4777
9.9788 | -93.4777
9.9791 | -93.4778
9.9794 | -93.478
9.9797 | -93.4781
9.98 | -93.4781
9.9803 | -93.4781
9.9806 | -93.4783
9.9809 | -93.4785
9.9812 | -93.4792
9.9815 | -93.4795
9.9818 | -93.4797
9.9822 | -93.4799
9.9825 | -93.4799
9.9828 | -93.48
9.9831 | -93.4802
9.9834 | -93.4802
9.9837 | -93.4803
9.984 | -93.4806
9.9843 | -93.4809
9.9846 | -93.4811
9.9849 | -93.4813
9.9852 | -93.4815
9.9855 | -93.4818
9.9859 | -93.4821
9.9862 | -93.4825
9.9865 | -93.4829
9.9868 | -93.4836
9.9871 | -93.4836
9.9874 | -93.4839
9.9877 | -93.4842
9.988 | -93.4842
9.9883 | -93.4845
9.9886 | -93.4848
9.9889 | -93.485
9.9893 | -93.4852
9.9896 | -93.4854
9.9899 | -93.4856
9.9902 | -93.4859
9.9905 | -93.4862
9.9908 | -93.4864
9.9911 | -93.4866
9.9914 | -93.4869
9.9917 | -93.4872
9.992 | -93.4874
9.9923 | -93.4876
9.9927 | -93.4879
9.993 | -93.4879
9.9933 | -93.4879
9.9936 | -93.4881
9.9939 | -93.4888
9.9942 | -93.4889
9.9945 | -93.4885
9.9948 | -93.4885
9.9951 | -93.4885
9.9954 | -93.4888
9.9957 | -93.4887
9.996 | -93.4882
9.9964 | -93.4876
9.9967 | -93.4871
9.997 | -93.4868
9.9973 | -93.4869
9.9976 | -93.4872
9.9979 | -93.4874
9.9982 | -93.4874
9.9985 | -93.4879
9.9988 | -93.4887
9.9991 | -93.4892
9.9994 | -93.4897
9.9998 | -93.4907
10.0001 | -93.4911
10.0004 | -93.4907
10.0007 | -93.4901
10.001 | -93.4901
10.0013 | -93.4906
10.0016 | -93.4907
10.0019 | -93.4906
10.0022 | -93.491
10.0025 | -93.4907
10.0028 | -93.4909
10.0031 | -93.4912
10.0035 | -93.4915
10.0038 | -93.4915
10.0041 | -93.4917
10.0044 | -93.4921
10.0047 | -93.492
10.005 | -93.4923
10.0053 | -93.4926
10.0056 | -93.493
10.0059 | -93.4921
10.0062 | -93.4919
10.0065 | -93.4919
10.0069 | -93.4924
10.0072 | -93.4925
10.0075 | -93.492
10.0078 | -93.4919
10.0081 | -93.492
10.0084 | -93.4919
10.0087 | -93.4922
10.009 | -93.4925
10.0093 | -93.4929
10.0096 | -93.4933
10.0099 | -93.4934
10.0103 | -93.4938
10.0106 | -93.4941
10.0109 | -93.494
10.0112 | -93.4941
10.0115 | -93.4943
10.0118 | -93.4944
10.0121 | -93.4943
10.0124 | -93.4951
10.0127 | -93.4961
10.013 | -93.4963
10.0133 | -93.4967
10.0136 | -93.4969
10.014 | -93.4969
10.0143 | -93.4974
10.0146 | -93.4981
10.0149 | -93.4985
10.0152 | -93.4986
10.0155 | -93.4983
10.0158 | -93.4983
10.0161 | -93.4984
10.0164 | -93.4983
10.0167 | -93.4983
10.017 | -93.4983
10.0174 | -93.4986
10.0177 | -93.4989
10.018 | -93.4991
10.0183 | -93.4991
10.0186 | -93.4992
10.0189 | -93.4995
10.0192 | -93.4997
10.0195 | -93.4993
10.0198 | -93.4998
10.0201 | -93.5001
10.0204 | -93.5
10.0207 | -93.5001
10.0211 | -93.5005
10.0214 | -93.5012
10.0217 | -93.5013
10.022 | -93.5016
10.0223 | -93.5021
10.0226 | -93.5024
10.0229 | -93.5026
10.0232 | -93.5026
10.0235 | -93.5024
10.0238 | -93.5025
10.0241 | -93.503
10.0245 | -93.5033
10.0248 | -93.5033
10.0251 | -93.5037
10.0254 | -93.5039
10.0257 | -93.5036
10.026 | -93.5035
10.0263 | -93.5038
10.0266 | -93.504
10.0269 | -93.5043
10.0272 | -93.5043
10.0275 | -93.5044
10.0279 | -93.5046
10.0282 | -93.5044
10.0285 | -93.5045
10.0288 | -93.5043
10.0291 | -93.5049
10.0294 | -93.5049
10.0297 | -93.5051
10.03 | -93.5055
10.0303 | -93.5056
10.0306 | -93.5066
10.0309 | -93.5066
10.0312 | -93.5064
10.0316 | -93.5067
10.0319 | -93.5067
10.0322 | -93.5064
10.0325 | -93.5067
10.0328 | -93.5073
10.0331 | -93.5077
10.0334 | -93.5077
10.0337 | -93.5078
10.034 | -93.5079
10.0343 | -93.5079
10.0346 | -93.5077
10.035 | -93.5077
10.0353 | -93.5082
10.0356 | -93.5084
10.0359 | -93.5084
10.0362 | -93.5085
10.0365 | -93.5084
10.0368 | -93.5082
10.0371 | -93.5082
10.0374 | -93.5086
10.0377 | -93.509
10.038 | -93.5092
10.0383 | -93.5092
10.0387 | -93.5092
10.039 | -93.51
10.0393 | -93.5101
10.0396 | -93.5104
10.0399 | -93.5102
10.0402 | -93.5108
10.0405 | -93.5109
10.0408 | -93.511
10.0411 | -93.5113
10.0414 | -93.5106
10.0417 | -93.5107
10.0421 | -93.5111
10.0424 | -93.5115
10.0427 | -93.512
10.043 | -93.512
10.0433 | -93.5121
10.0436 | -93.5121
10.0439 | -93.512
10.0442 | -93.512
10.0445 | -93.5122
10.0448 | -93.5123
10.0451 | -93.5127
10.0455 | -93.513
10.0458 | -93.5132
10.0461 | -93.513
10.0464 | -93.5131
10.0467 | -93.5127
10.047 | -93.5126
10.0473 | -93.5124
10.0476 | -93.5123
10.0479 | -93.5127
10.0482 | -93.5128
10.0485 | -93.5127
10.0488 | -93.5128
10.0492 | -93.5125
10.0495 | -93.5128
10.0498 | -93.5129
10.0501 | -93.5132
10.0504 | -93.5135
10.0507 | -93.5138
10.051 | -93.5141
10.0513 | -93.5145
10.0516 | -93.5151
10.0519 | -93.5156
10.0522 | -93.5154
10.0526 | -93.5154
10.0529 | -93.5154
10.0532 | -93.5157
10.0535 | -93.516
10.0538 | -93.5167
10.0541 | -93.5174
10.0544 | -93.5173
10.0547 | -93.5174
10.055 | -93.5174
10.0553 | -93.5176
10.0556 | -93.5176
10.0559 | -93.5177
10.0563 | -93.5178
10.0566 | -93.5176
10.0569 | -93.5172
10.0572 | -93.5169
10.0575 | -93.5164
10.0578 | -93.5156
10.0581 | -93.5158
10.0584 | -93.5159
10.0587 | -93.5157
10.059 | -93.5158
10.0593 | -93.5165
10.0597 | -93.5164
10.06 | -93.5163
10.0603 | -93.5158
10.0606 | -93.515
10.0609 | -93.5149
10.0612 | -93.515
10.0615 | -93.5151
10.0618 | -93.515
10.0621 | -93.5151
10.0624 | -93.5154
10.0627 | -93.5166
10.0631 | -93.5177
10.0634 | -93.5178
10.0637 | -93.5179
10.064 | -93.5183
10.0643 | -93.5184
10.0646 | -93.5185
10.0649 | -93.5182
10.0652 | -93.518
10.0655 | -93.5179
10.0658 | -93.5179
10.0661 | -93.5171
10.0664 | -93.5162
10.0668 | -93.5166
10.0671 | -93.5168
10.0674 | -93.5168
10.0677 | -93.5165
10.068 | -93.5159
10.0683 | -93.5152
10.0686 | -93.5147
10.0689 | -93.5143
10.0692 | -93.5141
10.0695 | -93.5142
10.0698 | -93.5144
10.0702 | -93.5143
10.0705 | -93.514
10.0708 | -93.5129
10.0711 | -93.5121
10.0714 | -93.512
10.0717 | -93.5121
10.072 | -93.512
10.0723 | -93.5123
10.0726 | -93.5123
10.0729 | -93.5122
10.0732 | -93.5121
10.0735 | -93.5121
10.0739 | -93.5121
10.0742 | -93.5122
10.0745 | -93.5121
10.0748 | -93.5117
10.0751 | -93.5118
10.0754 | -93.5121
10.0757 | -93.5126
10.076 | -93.5126
10.0763 | -93.513
10.0766 | -93.5132
10.0769 | -93.5133
10.0773 | -93.5134
10.0776 | -93.5137
10.0779 | -93.5137
10.0782 | -93.5141
10.0785 | -93.5146
10.0788 | -93.5148
10.0791 | -93.5152
10.0794 | -93.5152
10.0797 | -93.5155
10.08 | -93.5165
10.0803 | -93.5164
10.0807 | -93.5161
10.081 | -93.516
10.0813 | -93.5159
10.0816 | -93.5164
10.0819 | -93.5169
10.0822 | -93.5165
10.0825 | -93.5165
10.0828 | -93.517
10.0831 | -93.518
10.0834 | -93.5186
10.0837 | -93.5186
10.084 | -93.5185
10.0844 | -93.5188
10.0847 | -93.5188
10.085 | -93.519
10.0853 | -93.5192
10.0856 | -93.5194
10.0859 | -93.5189
10.0862 | -93.5188
10.0865 | -93.519
10.0868 | -93.5193
10.0871 | -93.5195
10.0874 | -93.5191
10.0878 | -93.5192
10.0881 | -93.5193
10.0884 | -93.5194
10.0887 | -93.5195
10.089 | -93.5194
10.0893 | -93.5191
10.0896 | -93.5189
10.0899 | -93.519
10.0902 | -93.5191
10.0905 | -93.5192
10.0908 | -93.5195
10.0911 | -93.5192
10.0915 | -93.5193
10.0918 | -93.5198
10.0921 | -93.5201
10.0924 | -93.5207
10.0927 | -93.5219
10.093 | -93.5221
10.0933 | -93.5224
10.0936 | -93.5225
10.0939 | -93.5224
10.0942 | -93.5224
10.0945 | -93.5225
10.0949 | -93.5228
10.0952 | -93.5229
10.0955 | -93.5227
10.0958 | -93.5224
10.0961 | -93.5225
10.0964 | -93.5224
10.0967 | -93.5224
10.097 | -93.5225
10.0973 | -93.5227
10.0976 | -93.5225
10.0979 | -93.5226
10.0983 | -93.5227