## CALIBRATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 69 |  | CTBTO
Transducer Block | 02-90C/P | 66 |  | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
Hood | 02-92B/P | 55
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE REFERENCE
Atmospheric pressure | 1012.4 | hPa |  | PR 000 150_-_CAL
Humidity | 19 | %/Hr |  | Measurement device
Internal Temperature of Calibration Bench | 23.14 | °C |  | IS CALIBRATOR
AUX Sensor pressure | 1017.5 | hPa |  | Sensor SoH at test start
AUX Sensor Temperature | 23.22 | °C |  | OK
SENSITIVITY AT 1 Hz
Output Parameters | Theoretical Value | Measured Value @1Hz | Margin (±) | Unit | Result
Sensitivity of Pressure Derivative output | 2 | 1.99 | 0.04 | mV/Pa/sec | OK | Mode d'emploi pour créer un constat de vérification
Sensitivity of Pressure output | 20 | 19.95 | 0.4 | mV/Pa | OK | Cliquer sur "DEMARRER LA SAISIE"
Sensitivity of Calibration coil | 6 | 5.96 | 0.13 | Pa/V | OK | Suivre les instructions affichées
Une fois la partie SENSOR IDENTIFICATION remplie, renseigner la partie ENVIRONNEMENTAL PARAMETERS (les données environnementales)
TRANSFER FUNCTION | [5 empty] | Renseigner la partie SENSITIVITY AT 1Hz :les sensibilités en voie de dérivée de Pression (Pressure Derivative), en voie de pression (Pressure) et de Calibration (calibration Coil)
Electrical |  | Acoustical |  |  |  | Une sauvegarde temporaire est toujours possible si le fichier doit être modifié (Fichier "enregistrer")
Vérifier toutes les données saisie notamment la partie CUSTOMER ("Client") et la partie VISA (Noms des personnes et date d'ajustage)
Cliquer sur "ENREGISTREMENT DEFINITIF" (Attention le fichier sera enregistré et verrouillé dans le dossier courant, il faut retirer manuellement la protection)
SENSOR SELF NOISE |  |  |  | Decision : considering the specifications given by the manufacturer, this sensor is
CONFORM
VISA
|  |  |  | Adjusted by | A.HUE
|  |  |  | on | 2016-03-15
|  |  |  | Approved by | A.HUE
|  |  |  | Qualification | Optronic Engineer
Self Noise at 1 HZ | 6.3533e-05 | Pa/√Hz |  | on | 2016-03-17 | -83.94 | (dBPa/√Hz)
| < 0,13 | mPa/√Hz
OK
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which define the verification report. It can be used to demonstrate the link between the measure equipment and the standard national or international equipment, provided it meets the recommendations of the documentation booklet X 07-015
[8 empty] | Client | Adresse Client
[8 empty] | CEA | C.E.A. Dam. Bruyères Domaine du Grand Rué 91297 ARPAJON CEDEX FRANCE
[8 empty] | CTBTO | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
[8 empty] | IPGP | Université Paris Diderot Paris 7 Bat Lamarck UFR STEP 35 rue H.Brion 6ème étage 75205 Paris CEDEX 13
[8 empty] | SANDIA | C/O SANDIATIONAL LABS ALBUQUERQUE, NM 87123 UNITED STATES
[8 empty] | ISLA | Haiwaï Institute of Geophysics & Planetology 73-984 Makako Bay Drive KAILUA-KONA HAWAÏ
ENREGISTREMENT DEFINITIF
DEMARRER LA SAISIE

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.1331 | 12.0175
0.097656 | -34.0786 | 7.8786
0.12207 | -33.9876 | 5.4557
0.14648 | -33.9627 | 3.8484
0.1709 | -33.9201 | 3.0799
0.19531 | -33.9592 | 2.5919
0.21973 | -33.9788 | 2.4392
0.24414 | -33.932 | 1.7152
0.26855 | -33.9673 | 1.7176
0.29297 | -34.0123 | 0.81525
0.31738 | -33.9561 | 0.91498
0.3418 | -33.9476 | 0.6194
0.36621 | -33.9957 | 0.57021
0.39062 | -33.9463 | 0.4458
0.41504 | -34.015 | 0.14417
0.43945 | -33.9649 | 0.16579
0.46387 | -33.9391 | -0.46379
0.48828 | -33.9071 | 0.17109
0.5127 | -33.9868 | -0.49229
0.53711 | -33.9202 | -0.21134
0.56152 | -33.9941 | -0.671
0.58594 | -34.005 | -0.96309
0.61035 | -33.9957 | -0.64751
0.63477 | -34.012 | -0.51133
0.65918 | -34.0152 | -0.87212
0.68359 | -33.937 | -1.0182
0.70801 | -33.9892 | -0.89653
0.73242 | -34.0188 | -0.94023
0.75684 | -33.9916 | -1.2998
0.78125 | -34.0057 | -1.1634
0.80566 | -34.0356 | -1.1256
0.83008 | -33.9862 | -1.4587
0.85449 | -34.0009 | -1.9306
0.87891 | -33.9864 | -1.6733
0.90332 | -33.9915 | -1.7793
0.92773 | -34.0187 | -1.368
0.95215 | -34.0365 | -1.8539
0.97656 | -33.9991 | -1.8059
1.001 | -34.0253 | -2.1818
1.0254 | -34.069 | -2.0511
1.0498 | -33.9582 | -2.1836
1.0742 | -34.0134 | -2.8031
1.0986 | -34.0395 | -2.0781
1.123 | -33.9698 | -2.5524
1.1475 | -34.0546 | -2.8548
1.1719 | -34.0597 | -2.9
1.1963 | -34.0562 | -2.4947
1.2207 | -34.121 | -2.4483
1.2451 | -34.0129 | -2.2844
1.2695 | -34.0022 | -3.0475
1.2939 | -33.9923 | -2.92
1.3184 | -33.9981 | -2.9274
1.3428 | -34.0454 | -3.0939
1.3672 | -34.0362 | -3.0626
1.3916 | -34.0298 | -3.3022
1.416 | -33.9938 | -2.798
1.4404 | -34.0028 | -3.3064
1.4648 | -33.9684 | -3.2955
1.4893 | -33.9611 | -3.5796
1.5137 | -34.0096 | -3.5408
1.5381 | -33.9864 | -3.746
1.5625 | -34.039 | -3.4098
1.5869 | -34.0523 | -3.714
1.6113 | -34.1216 | -3.6503
1.6357 | -33.9885 | -4.1048
1.6602 | -34.0224 | -3.9247
1.6846 | -33.9645 | -4.2865
1.709 | -33.9902 | -4.1601
1.7334 | -33.9929 | -4.2664
1.7578 | -34.0944 | -4.3048
1.7822 | -33.9685 | -4.0889
1.8066 | -34.0228 | -4.4816
1.8311 | -34.028 | -4.5472
1.8555 | -34.1113 | -4.4377
1.8799 | -34.0178 | -4.2298
1.9043 | -34.028 | -4.864
1.9287 | -34.0216 | -4.7238
1.9531 | -34.0131 | -4.9767
1.9775 | -34.0371 | -4.2457
2.002 | -34.0331 | -5.1901
2.0264 | -34.0302 | -5.0765
2.0508 | -34.0998 | -4.669
2.0752 | -34.0671 | -4.6519
2.0996 | -34.11 | -5.2181
2.124 | -34.1091 | -5.0131
2.1484 | -34.0382 | -5.1865
2.1729 | -34.0211 | -5.1686
2.1973 | -33.9992 | -5.381
2.2217 | -33.9803 | -5.1437
2.2461 | -34.0014 | -5.6801
2.2705 | -34.0294 | -5.5074
2.2949 | -34.0292 | -5.4258
2.3193 | -34.0553 | -6.1066
2.3438 | -34.002 | -5.7671
2.3682 | -34.1016 | -5.8885
2.3926 | -33.9908 | -5.6269
2.417 | -34.0744 | -6.1269
2.4414 | -34.0569 | -5.3782
2.4658 | -34.0659 | -5.8135
2.4902 | -34.0738 | -5.9605
2.5146 | -34.0608 | -5.8746
2.5391 | -34.084 | -6.0466
2.5635 | -33.9801 | -5.9516
2.5879 | -34.0381 | -6.3615
2.6123 | -34.0509 | -6.4317
2.6367 | -34.0312 | -6.4835
2.6611 | -34.0888 | -6.1732
2.6855 | -34.0468 | -6.5037
2.71 | -34.0344 | -6.8067
2.7344 | -34.0782 | -6.2508
2.7588 | -34.0244 | -6.0223
2.7832 | -34.0619 | -6.9333
2.8076 | -34.0477 | -6.2779
2.832 | -34.0424 | -6.7592
2.8564 | -34.0341 | -6.8832
2.8809 | -34.0875 | -7.1687
2.9053 | -34.0939 | -7.3274
2.9297 | -34.0717 | -7.0445
2.9541 | -34.0783 | -6.9733
2.9785 | -34.0016 | -7.3288
3.0029 | -34.0819 | -6.9582
3.0273 | -34.0325 | -7.7254
3.0518 | -34.0578 | -7.2126
3.0762 | -34.0731 | -7.6479
3.1006 | -34.0465 | -7.2254
3.125 | -34.0596 | -7.7417
3.1494 | -34.0308 | -7.7292
3.1738 | -34.0914 | -7.639
3.1982 | -34.1316 | -8.0546
3.2227 | -34.0818 | -7.9371
3.2471 | -34.1217 | -7.836
3.2715 | -34.0259 | -7.3174
3.2959 | -34.0813 | -7.5291
3.3203 | -34.0766 | -7.6675
3.3447 | -34.1039 | -7.9954
3.3691 | -34.1177 | -8.2126
3.3936 | -34.0876 | -8.4253
3.418 | -34.0866 | -7.9622
3.4424 | -34.0881 | -8.404
3.4668 | -34.1146 | -8.001
3.4912 | -34.1176 | -8.2457
3.5156 | -34.1282 | -8.1382
3.54 | -34.0187 | -8.0966
3.5645 | -34.0851 | -8.8722
3.5889 | -34.0496 | -9.3028
3.6133 | -34.1501 | -9.1973
3.6377 | -34.0717 | -9.2771
3.6621 | -34.0933 | -9.4792
3.6865 | -34.088 | -8.9713
3.7109 | -34.1648 | -8.9839
3.7354 | -34.0887 | -8.6842
3.7598 | -34.0947 | -9.6907
3.7842 | -34.115 | -9.3687
3.8086 | -34.0886 | -9.3781
3.833 | -34.1169 | -8.9197
3.8574 | -34.0961 | -9.3257
3.8818 | -34.1524 | -9.1973
3.9062 | -34.0906 | -9.223
3.9307 | -34.1064 | -9.2905
3.9551 | -34.0889 | -9.5067
3.9795 | -34.1621 | -9.8581
4.0039 | -34.1019 | -9.7587
4.0283 | -34.1283 | -10.0725
4.0527 | -34.1027 | -10.1028
4.0771 | -34.2087 | -10.0869
4.1016 | -34.1579 | -9.8174
4.126 | -34.1463 | -10.2533
4.1504 | -34.1029 | -9.7664
4.1748 | -34.1135 | -10.0663
4.1992 | -34.0655 | -10.2993
4.2236 | -34.1137 | -10.5405
4.248 | -34.0725 | -9.9562
4.2725 | -34.112 | -10.673
4.2969 | -34.1913 | -10.4445
4.3213 | -34.1389 | -10.3691
4.3457 | -34.1152 | -10.696
4.3701 | -34.1424 | -10.6359
4.3945 | -34.1541 | -10.4773
4.4189 | -34.1257 | -10.2973
4.4434 | -34.0916 | -10.5569
4.4678 | -34.1271 | -10.682
4.4922 | -34.1642 | -10.7727
4.5166 | -34.1049 | -10.7998
4.541 | -34.1213 | -10.9272
4.5654 | -34.0564 | -11.017
4.5898 | -34.1723 | -11.6442
4.6143 | -34.0561 | -10.994
4.6387 | -34.1317 | -11.2823
4.6631 | -34.1753 | -11.3419
4.6875 | -34.1879 | -11.4772
4.7119 | -34.1435 | -11.0322
4.7363 | -34.1091 | -11.2551
4.7607 | -34.152 | -11.3631
4.7852 | -34.0877 | -12.1101
4.8096 | -34.1954 | -11.4799
4.834 | -34.1521 | -11.8933
4.8584 | -34.1385 | -11.66
4.8828 | -34.1721 | -11.4272
4.9072 | -34.1911 | -11.9584
4.9316 | -34.127 | -12.0119
4.9561 | -34.1898 | -11.9245
4.9805 | -34.1131 | -12.1299
5.0049 | -34.1742 | -12.2958
5.0293 | -34.1098 | -12.317
5.0537 | -34.2015 | -11.9553
5.0781 | -34.0926 | -12.0166
5.1025 | -34.1705 | -11.9805
5.127 | -34.1465 | -12.3469
5.1514 | -34.1355 | -12.5652
5.1758 | -34.1559 | -12.2571
5.2002 | -34.0528 | -12.6775
5.2246 | -34.0924 | -12.8731
5.249 | -34.2229 | -13.1801
5.2734 | -34.1496 | -12.3029
5.2979 | -34.2375 | -12.8574
5.3223 | -34.2066 | -12.6442
5.3467 | -34.1967 | -12.9382
5.3711 | -34.1744 | -13.1289
5.3955 | -34.174 | -13.3142
5.4199 | -34.1816 | -13.2582
5.4443 | -34.1634 | -13.103
5.4688 | -34.2279 | -13.5537
5.4932 | -34.1685 | -12.7525
5.5176 | -34.1464 | -13.2113
5.542 | -34.1853 | -13.3925
5.5664 | -34.1603 | -13.0022
5.5908 | -34.1733 | -13.2739
5.6152 | -34.1912 | -14.0328
5.6396 | -34.1261 | -13.4572
5.6641 | -34.1133 | -13.6752
5.6885 | -34.2011 | -13.5975
5.7129 | -34.1102 | -13.7473
5.7373 | -34.1458 | -13.6058
5.7617 | -34.1606 | -13.5389
5.7861 | -34.1808 | -13.6005
5.8105 | -34.154 | -14.3024
5.835 | -34.2107 | -13.7976
5.8594 | -34.2255 | -14.2357
5.8838 | -34.2419 | -14.1694
5.9082 | -34.1474 | -13.9586
5.9326 | -34.1855 | -14.5836
5.957 | -34.2012 | -14.2836
5.9814 | -34.2034 | -14.4216
6.0059 | -34.1945 | -14.3599
6.0303 | -34.201 | -14.3479
6.0547 | -34.1326 | -14.6874
6.0791 | -34.1456 | -14.512
6.1035 | -34.2308 | -14.5925
6.1279 | -34.1626 | -14.8408
6.1523 | -34.3073 | -15.1651
6.1768 | -34.1738 | -14.7357
6.2012 | -34.2143 | -14.7041
6.2256 | -34.2282 | -15.349
6.25 | -34.1593 | -15.0008
6.2744 | -34.1666 | -15.0259
6.2988 | -34.1927 | -15.0865
6.3232 | -34.1159 | -14.974
6.3477 | -34.2159 | -15.5263
6.3721 | -34.2 | -15.156
6.3965 | -34.2352 | -15.4919
6.4209 | -34.1804 | -15.1846
6.4453 | -34.1612 | -15.7713
6.4697 | -34.236 | -15.3185
6.4941 | -34.28 | -15.8691
6.5186 | -34.2372 | -15.2636
6.543 | -34.2714 | -15.625
6.5674 | -34.2454 | -16.0709
6.5918 | -34.1965 | -16.1828
6.6162 | -34.2189 | -15.7869
6.6406 | -34.2104 | -15.8776
6.665 | -34.1779 | -15.7804
6.6895 | -34.2052 | -15.8997
6.7139 | -34.287 | -15.907
6.7383 | -34.1497 | -16.4529
6.7627 | -34.208 | -16.4692
6.7871 | -34.2872 | -16.0422
6.8115 | -34.2948 | -16.3786
6.8359 | -34.1495 | -16.4882
6.8604 | -34.2619 | -16.9942
6.8848 | -34.2641 | -16.3182
6.9092 | -34.3017 | -16.3212
6.9336 | -34.2997 | -16.5127
6.958 | -34.2562 | -16.8884
6.9824 | -34.2293 | -16.1894
7.0068 | -34.2243 | -16.746
7.0312 | -34.2553 | -16.9322
7.0557 | -34.2103 | -16.5052
7.0801 | -34.2884 | -17.1045
7.1045 | -34.2615 | -16.9867
7.1289 | -34.3228 | -17.2978
7.1533 | -34.2808 | -17.5022
7.1777 | -34.2601 | -16.6734
7.2021 | -34.2389 | -17.0835
7.2266 | -34.2942 | -17.4571
7.251 | -34.2442 | -17.4917
7.2754 | -34.2937 | -17.5451
7.2998 | -34.2217 | -17.759
7.3242 | -34.2766 | -17.561
7.3486 | -34.2573 | -17.2951
7.373 | -34.2564 | -17.7879
7.3975 | -34.2178 | -17.8829
7.4219 | -34.2852 | -17.965
7.4463 | -34.3134 | -17.6939
7.4707 | -34.2302 | -18.1803
7.4951 | -34.3272 | -17.8135
7.5195 | -34.2905 | -18.226
7.5439 | -34.2423 | -17.7603
7.5684 | -34.2754 | -18.0856
7.5928 | -34.3414 | -18.0488
7.6172 | -34.2669 | -18.0937
7.6416 | -34.2644 | -18.5174
7.666 | -34.2727 | -18.1349
7.6904 | -34.306 | -18.1853
7.7148 | -34.2939 | -18.4624
7.7393 | -34.3506 | -18.3551
7.7637 | -34.328 | -18.8234
7.7881 | -34.3231 | -18.7501
7.8125 | -34.2754 | -18.6301
7.8369 | -34.2714 | -18.8721
7.8613 | -34.266 | -18.9292
7.8857 | -34.3053 | -18.6711
7.9102 | -34.3611 | -18.5061
7.9346 | -34.2858 | -18.8342
7.959 | -34.3309 | -18.6475
7.9834 | -34.286 | -19.0569
8.0078 | -34.3199 | -19.2195
8.0322 | -34.298 | -19.2104
8.0566 | -34.3197 | -18.7597
8.0811 | -34.2543 | -19.0608
8.1055 | -34.3315 | -19.8277
8.1299 | -34.2634 | -19.435
8.1543 | -34.3332 | -19.3089
8.1787 | -34.2939 | -19.4536
8.2031 | -34.3639 | -19.5058
8.2275 | -34.2822 | -19.1789
8.252 | -34.3629 | -19.9916
8.2764 | -34.3153 | -19.6978
8.3008 | -34.3376 | -19.3822
8.3252 | -34.3539 | -19.5636
8.3496 | -34.3596 | -19.6903
8.374 | -34.3342 | -20.1143
8.3984 | -34.2818 | -19.9876
8.4229 | -34.3589 | -19.6184
8.4473 | -34.2721 | -19.9931
8.4717 | -34.3007 | -19.7933
8.4961 | -34.2884 | -20.0144
8.5205 | -34.3121 | -20.1491
8.5449 | -34.2873 | -20.4631
8.5693 | -34.351 | -20.2396
8.5938 | -34.3004 | -20.3749
8.6182 | -34.3641 | -20.7279
8.6426 | -34.3764 | -20.403
8.667 | -34.3634 | -20.6217
8.6914 | -34.2904 | -20.2306
8.7158 | -34.3481 | -20.5849
8.7402 | -34.3501 | -20.6073
8.7646 | -34.3008 | -20.9469
8.7891 | -34.3705 | -20.8151
8.8135 | -34.3034 | -20.3266
8.8379 | -34.3548 | -20.5796
8.8623 | -34.3644 | -21.1635
8.8867 | -34.4358 | -20.846
8.9111 | -34.3988 | -21.0535
8.9355 | -34.4003 | -20.9764
8.96 | -34.3591 | -21.0703
8.9844 | -34.4215 | -21.7439
9.0088 | -34.3543 | -21.595
9.0332 | -34.4159 | -21.0786
9.0576 | -34.3668 | -21.1877
9.082 | -34.3429 | -21.5666
9.1064 | -34.3731 | -21.0471
9.1309 | -34.4014 | -21.5632
9.1553 | -34.4239 | -21.9152
9.1797 | -34.3472 | -21.6019
9.2041 | -34.3713 | -21.5412
9.2285 | -34.3901 | -22.0413
9.2529 | -34.3483 | -21.6002
9.2773 | -34.3394 | -21.9723
9.3018 | -34.3958 | -21.8132
9.3262 | -34.3356 | -21.9933
9.3506 | -34.4323 | -22.124
9.375 | -34.3809 | -22.0753
9.3994 | -34.4504 | -22.2831
9.4238 | -34.4016 | -22.3938
9.4482 | -34.455 | -22.2769
9.4727 | -34.3966 | -21.8864
9.4971 | -34.4242 | -22.7562
9.5215 | -34.3771 | -22.4618
9.5459 | -34.4706 | -22.8651
9.5703 | -34.4468 | -22.2813
9.5947 | -34.411 | -22.7586
9.6191 | -34.3585 | -22.3908
9.6436 | -34.4398 | -22.9334
9.668 | -34.434 | -22.6977
9.6924 | -34.4267 | -22.865
9.7168 | -34.4211 | -22.8725
9.7412 | -34.4306 | -23.1021
9.7656 | -34.4114 | -23.0935
9.79 | -34.4326 | -23.0785
9.8145 | -34.4616 | -22.89
9.8389 | -34.361 | -23.4968
9.8633 | -34.4399 | -23.619
9.8877 | -34.4636 | -23.3665
9.9121 | -34.3959 | -23.3623
9.9365 | -34.4092 | -23.4338
9.9609 | -34.4236 | -23.2556
9.9854 | -34.3959 | -23.2842
10.0098 | -34.4246 | -23.4717

## ACOUS_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.7256 | 8.7147
0.097656 | -34.2491 | 2.3277
0.12207 | -33.915 | 2.9246
0.14648 | -34.0986 | 1.7374
0.1709 | -34.008 | 1.2351
0.19531 | -34.0687 | 0.11635
0.21973 | -34.0532 | 0.49791
0.24414 | -34.0892 | -0.82343
0.26855 | -34.0849 | -0.099
0.29297 | -34.0297 | -1.3943
0.31738 | -34.0737 | -0.86377
0.3418 | -34.0123 | -1.7222
0.36621 | -34.0964 | -1.3216
0.39062 | -34.0009 | -2.0716
0.41504 | -34.1106 | -1.6119
0.43945 | -34.0125 | -2.0438
0.46387 | -34.1137 | -2.2163
0.48828 | -33.9875 | -2.2613
0.5127 | -34.1342 | -2.5527
0.53711 | -34.0235 | -2.251
0.56152 | -34.1159 | -2.8913
0.58594 | -34.0281 | -2.4411
0.61035 | -34.104 | -3.241
0.63477 | -34.0567 | -2.4986
0.65918 | -34.0821 | -3.3907
0.68359 | -34.0511 | -2.6341
0.70801 | -34.0511 | -3.4709
0.73242 | -34.0934 | -2.8257
0.75684 | -34.053 | -3.6573
0.78125 | -34.1107 | -2.977
0.80566 | -34.0318 | -3.7547
0.83008 | -34.1296 | -3.3262
0.85449 | -34.0081 | -3.7803
0.87891 | -34.0998 | -3.4761
0.90332 | -33.9954 | -3.5356
0.92773 | -34.1415 | -3.5417
0.95215 | -34.0343 | -3.5475
0.97656 | -34.1453 | -4.02
1.001 | -34.0361 | -3.6609
1.0254 | -34.1393 | -4.2108
1.0498 | -34.081 | -3.7055
1.0742 | -34.1328 | -4.5509
1.0986 | -34.0627 | -3.8972
1.123 | -34.0866 | -4.5841
1.1475 | -34.0896 | -3.7911
1.1719 | -34.0814 | -4.7151
1.1963 | -34.1113 | -3.8509
1.2207 | -34.0684 | -4.8251
1.2451 | -34.1475 | -4.0827
1.2695 | -34.0464 | -4.7285
1.2939 | -34.1533 | -4.2158
1.3184 | -34.016 | -4.621
1.3428 | -34.1791 | -4.444
1.3672 | -34.051 | -4.6448
1.3916 | -34.1553 | -4.7794
1.416 | -34.0175 | -4.6949
1.4404 | -34.1458 | -4.9062
1.4648 | -34.0284 | -4.3816
1.4893 | -34.1526 | -4.9349
1.5137 | -34.0774 | -4.4328
1.5381 | -34.1483 | -5.3404
1.5625 | -34.1084 | -4.4943
1.5869 | -34.1322 | -5.4092
1.6113 | -34.1314 | -4.7483
1.6357 | -34.0828 | -5.6873
1.6602 | -34.1233 | -4.9061
1.6846 | -34.0676 | -5.4893
1.709 | -34.1519 | -4.9741
1.7334 | -34.038 | -5.5456
1.7578 | -34.1803 | -5.214
1.7822 | -34.0371 | -5.3757
1.8066 | -34.1759 | -5.3103
1.8311 | -34.0465 | -5.271
1.8555 | -34.2011 | -5.5248
1.8799 | -34.0697 | -5.3202
1.9043 | -34.1689 | -5.9262
1.9287 | -34.0493 | -5.3613
1.9531 | -34.1443 | -6.0039
1.9775 | -34.0867 | -5.1271
2.002 | -34.1403 | -6.1236
2.0264 | -34.1051 | -5.3084
2.0508 | -34.1128 | -6.0429
2.0752 | -34.1474 | -5.3776
2.0996 | -34.1214 | -6.2122
2.124 | -34.1752 | -5.5401
2.1484 | -34.0964 | -6.4356
2.1729 | -34.1747 | -5.845
2.1973 | -34.0796 | -6.1856
2.2217 | -34.1819 | -5.9668
2.2461 | -34.0795 | -6.1895
2.2705 | -34.2228 | -6.1658
2.2949 | -34.0927 | -6.223
2.3193 | -34.1886 | -6.3658
2.3438 | -34.0916 | -6.0699
2.3682 | -34.1626 | -6.5571
2.3926 | -34.083 | -5.8951
2.417 | -34.1506 | -6.6587
2.4414 | -34.1211 | -5.8623
2.4658 | -34.1817 | -6.7494
2.4902 | -34.1658 | -5.9669
2.5146 | -34.1426 | -6.983
2.5391 | -34.1586 | -6.2324
2.5635 | -34.1228 | -6.9276
2.5879 | -34.1851 | -6.3268
2.6123 | -34.0993 | -6.8247
2.6367 | -34.189 | -6.4576
2.6611 | -34.0982 | -6.9222
2.6855 | -34.22 | -6.6341
2.71 | -34.0747 | -6.845
2.7344 | -34.2269 | -6.871
2.7588 | -34.1065 | -6.5858
2.7832 | -34.1981 | -7.1481
2.8076 | -34.0908 | -6.6299
2.832 | -34.1997 | -7.1831
2.8564 | -34.1461 | -6.7599
2.8809 | -34.1737 | -7.5992
2.9053 | -34.1289 | -6.7283
2.9297 | -34.1638 | -7.577
2.9541 | -34.1805 | -6.8778
2.9785 | -34.1302 | -7.6193
3.0029 | -34.1874 | -6.9035
3.0273 | -34.1284 | -7.6383
3.0518 | -34.2247 | -6.9538
3.0762 | -34.1192 | -7.5273
3.1006 | -34.2207 | -7.1887
3.125 | -34.1139 | -7.5044
3.1494 | -34.2291 | -7.4326
3.1738 | -34.1154 | -7.5747
3.1982 | -34.2332 | -7.6858
3.2227 | -34.0922 | -7.4274
3.2471 | -34.2109 | -7.8748
3.2715 | -34.1329 | -7.367
3.2959 | -34.2174 | -7.8335
3.3203 | -34.1404 | -7.2442
3.3447 | -34.1849 | -8.0955
3.3691 | -34.1805 | -7.3461
3.3936 | -34.1826 | -8.0757
3.418 | -34.1944 | -7.4291
3.4424 | -34.1659 | -8.3195
3.4668 | -34.2329 | -7.5869
3.4912 | -34.1502 | -8.3832
3.5156 | -34.2629 | -7.8361
3.54 | -34.1127 | -8.2605
3.5645 | -34.2472 | -7.9951
3.5889 | -34.1454 | -8.192
3.6133 | -34.2594 | -8.128
3.6377 | -34.1237 | -8.2135
3.6621 | -34.2401 | -8.3921
3.6865 | -34.1365 | -7.9486
3.7109 | -34.2578 | -8.4319
3.7354 | -34.1789 | -8.0607
3.7598 | -34.2304 | -8.6803
3.7842 | -34.1845 | -8.1636
3.8086 | -34.2025 | -8.7625
3.833 | -34.1871 | -8.0833
3.8574 | -34.1913 | -8.7729
3.8818 | -34.2418 | -8.1589
3.9062 | -34.18 | -8.9426
3.9307 | -34.2341 | -8.1664
3.9551 | -34.1724 | -8.8533
3.9795 | -34.2921 | -8.4949
4.0039 | -34.1565 | -8.8842
4.0283 | -34.2689 | -8.6225
4.0527 | -34.134 | -8.8612
4.0771 | -34.2819 | -8.9229
4.1016 | -34.1833 | -8.7188
4.126 | -34.2828 | -9.0419
4.1504 | -34.1734 | -8.7788
4.1748 | -34.2681 | -9.357
4.1992 | -34.195 | -8.7449
4.2236 | -34.2303 | -9.5222
4.248 | -34.1902 | -8.6857
4.2725 | -34.207 | -9.4202
4.2969 | -34.2125 | -8.5978
4.3213 | -34.2298 | -9.493
4.3457 | -34.2383 | -8.8369
4.3701 | -34.1821 | -9.557
4.3945 | -34.2651 | -8.9329
4.4189 | -34.1831 | -9.4713
4.4434 | -34.3068 | -9.003
4.4678 | -34.2091 | -9.5194
4.4922 | -34.2926 | -9.5161
4.5166 | -34.1541 | -9.4796
4.541 | -34.3011 | -9.5797
4.5654 | -34.175 | -9.3207
4.5898 | -34.2942 | -9.833
4.6143 | -34.1878 | -9.2297
4.6387 | -34.2718 | -9.9543
4.6631 | -34.2175 | -9.2991
4.6875 | -34.2604 | -10.0651
4.7119 | -34.223 | -9.3404
4.7363 | -34.2423 | -9.8554
4.7607 | -34.2538 | -9.3631
4.7852 | -34.2417 | -10.1825
4.8096 | -34.2837 | -9.5857
4.834 | -34.2061 | -10.2127
4.8584 | -34.3215 | -9.5147
4.8828 | -34.1901 | -9.9303
4.9072 | -34.3251 | -9.7385
4.9316 | -34.1965 | -10.1878
4.9561 | -34.3246 | -10.1235
4.9805 | -34.1856 | -9.9334
5.0049 | -34.3004 | -10.3541
5.0293 | -34.1692 | -9.9193
5.0537 | -34.2837 | -10.4584
5.0781 | -34.2024 | -9.7677
5.1025 | -34.309 | -10.6221
5.127 | -34.2313 | -9.8503
5.1514 | -34.2912 | -10.6776
5.1758 | -34.2683 | -9.9358
5.2002 | -34.2797 | -10.7776
5.2246 | -34.2864 | -9.951
5.249 | -34.2179 | -10.7182
5.2734 | -34.3019 | -9.9875
5.2979 | -34.2393 | -10.8376
5.3223 | -34.347 | -10.3701
5.3467 | -34.2095 | -10.6607
5.3711 | -34.3379 | -10.4673
5.3955 | -34.1936 | -10.6855
5.4199 | -34.3657 | -10.8059
5.4443 | -34.2228 | -10.5951
5.4688 | -34.3177 | -11.0783
5.4932 | -34.2003 | -10.5193
5.5176 | -34.3138 | -11.0411
5.542 | -34.2382 | -10.5627
5.5664 | -34.3101 | -11.2632
5.5908 | -34.2568 | -10.568
5.6152 | -34.2937 | -11.4317
5.6396 | -34.2653 | -10.7575
5.6641 | -34.2522 | -11.4602
5.6885 | -34.3041 | -10.6972
5.7129 | -34.231 | -11.457
5.7373 | -34.3105 | -10.8188
5.7617 | -34.2528 | -11.3031
5.7861 | -34.3795 | -10.7121
5.8105 | -34.2298 | -11.2924
5.835 | -34.396 | -11.1808
5.8594 | -34.224 | -11.2683
5.8838 | -34.3693 | -11.3267
5.9082 | -34.2357 | -11.3053
5.9326 | -34.3633 | -11.6971
5.957 | -34.2589 | -11.0067
5.9814 | -34.3504 | -11.6951
6.0059 | -34.2708 | -11.1729
6.0303 | -34.3177 | -11.9035
6.0547 | -34.2804 | -11.2157
6.0791 | -34.2916 | -11.769
6.1035 | -34.3123 | -10.9626
6.1279 | -34.2781 | -12.0485
6.1523 | -34.3387 | -11.3157
6.1768 | -34.2497 | -12.1459
6.2012 | -34.3485 | -11.5671
6.2256 | -34.2548 | -11.8511
6.25 | -34.3822 | -11.5011
6.2744 | -34.2701 | -11.7652
6.2988 | -34.37 | -11.9105
6.3232 | -34.2638 | -11.6485
6.3477 | -34.3825 | -11.7455
6.3721 | -34.2484 | -11.5554
6.3965 | -34.3673 | -12.1244
6.4209 | -34.2714 | -11.6924
6.4453 | -34.3519 | -12.2184
6.4697 | -34.2984 | -11.5863
6.4941 | -34.3518 | -12.5776
6.5186 | -34.3446 | -11.6461
6.543 | -34.3424 | -12.6817
6.5674 | -34.3675 | -11.7907
6.5918 | -34.3205 | -12.7807
6.6162 | -34.3654 | -12.0157
6.6406 | -34.259 | -12.569
6.665 | -34.412 | -12.0759
6.6895 | -34.2772 | -12.3784
6.7139 | -34.3978 | -12.2855
6.7383 | -34.2643 | -12.4511
6.7627 | -34.3796 | -12.5089
6.7871 | -34.2395 | -12.2993
6.8115 | -34.4209 | -12.623
6.8359 | -34.313 | -12.4881
6.8604 | -34.4026 | -12.8641
6.8848 | -34.2881 | -12.2516
6.9092 | -34.3832 | -12.926
6.9336 | -34.3575 | -12.4172
6.958 | -34.3551 | -13.1415
6.9824 | -34.3538 | -12.2914
7.0068 | -34.3457 | -13.0796
7.0312 | -34.3737 | -12.2869
7.0557 | -34.3406 | -13.0665
7.0801 | -34.3953 | -12.3855
7.1045 | -34.3418 | -12.9663
7.1289 | -34.4505 | -12.7055
7.1533 | -34.3131 | -13.0801
7.1777 | -34.4378 | -12.8587
7.2021 | -34.3363 | -13.0397
7.2266 | -34.4707 | -12.9917
7.251 | -34.3621 | -13.105
7.2754 | -34.4369 | -13.2133
7.2998 | -34.3006 | -12.6387
7.3242 | -34.3909 | -13.4208
7.3486 | -34.3642 | -12.7994
7.373 | -34.4023 | -13.4886
7.3975 | -34.3776 | -12.7254
7.4219 | -34.3983 | -13.7292
7.4463 | -34.3924 | -12.8823
7.4707 | -34.3573 | -13.6222
7.4951 | -34.3736 | -13.0493
7.5195 | -34.3517 | -13.5832
7.5439 | -34.4218 | -13.1593
7.5684 | -34.3159 | -13.5008
7.5928 | -34.4104 | -13.2743
7.6172 | -34.3068 | -13.4292
7.6416 | -34.4738 | -13.4004
7.666 | -34.3272 | -13.1364
7.6904 | -34.472 | -13.7867
7.7148 | -34.3609 | -13.3698
7.7393 | -34.4508 | -13.8705
7.7637 | -34.3991 | -13.4181
7.7881 | -34.4635 | -14.193
7.8125 | -34.3847 | -13.4619
7.8369 | -34.4007 | -14.2722
7.8613 | -34.4005 | -13.3107
7.8857 | -34.3846 | -14.515
7.9102 | -34.4592 | -13.4476
7.9346 | -34.3606 | -14.3521
7.959 | -34.4635 | -13.617
7.9834 | -34.365 | -14.3158
8.0078 | -34.4927 | -13.6718
8.0322 | -34.4011 | -14.222
8.0566 | -34.5036 | -13.9181
8.0811 | -34.3769 | -13.9138
8.1055 | -34.4787 | -13.9369
8.1299 | -34.3486 | -14.0069
8.1543 | -34.4784 | -14.4228
8.1787 | -34.3799 | -13.8659
8.2031 | -34.4304 | -14.2448
8.2275 | -34.4012 | -13.874
8.252 | -34.4603 | -14.6022
8.2764 | -34.43 | -14.095
8.3008 | -34.4475 | -14.6335
8.3252 | -34.4602 | -14.0654
8.3496 | -34.4195 | -14.8465
8.374 | -34.5005 | -13.9968
8.3984 | -34.398 | -15.0976
8.4229 | -34.4622 | -14.3352
8.4473 | -34.421 | -14.7863
8.4717 | -34.4733 | -14.2672
8.4961 | -34.3711 | -14.8794
8.5205 | -34.5169 | -14.2838
8.5449 | -34.3878 | -14.7131
8.5693 | -34.5403 | -14.6864
8.5938 | -34.3804 | -14.3372
8.6182 | -34.5236 | -15.0218
8.6426 | -34.4043 | -14.3832
8.667 | -34.4827 | -15.0872
8.6914 | -34.4098 | -14.1802
8.7158 | -34.4537 | -15.3989
8.7402 | -34.429 | -14.4204
8.7646 | -34.4545 | -15.1346
8.7891 | -34.4945 | -14.5183
8.8135 | -34.447 | -15.4796
8.8379 | -34.5382 | -14.379
8.8623 | -34.452 | -15.0341
8.8867 | -34.5599 | -14.8314
8.9111 | -34.4281 | -15.0489
8.9355 | -34.55 | -15.0763
8.96 | -34.442 | -15.1607
8.9844 | -34.5625 | -15.2584
9.0088 | -34.4396 | -15.3162
9.0332 | -34.5444 | -15.3012
9.0576 | -34.3766 | -15.051
9.082 | -34.579 | -15.386
9.1064 | -34.4738 | -14.779
9.1309 | -34.5459 | -15.726
9.1553 | -34.4651 | -14.9154
9.1797 | -34.4721 | -15.8771
9.2041 | -34.4894 | -15.0527
9.2285 | -34.476 | -15.7329
9.2529 | -34.4879 | -15.0241
9.2773 | -34.4405 | -15.9774
9.3018 | -34.5416 | -14.9571
9.3262 | -34.4711 | -15.8134
9.3506 | -34.595 | -15.3628
9.375 | -34.442 | -15.5371
9.3994 | -34.5468 | -15.2629
9.4238 | -34.4689 | -15.4336
9.4482 | -34.6138 | -15.3998
9.4727 | -34.4621 | -15.5309
9.4971 | -34.579 | -16.0826
9.5215 | -34.4535 | -15.5421
9.5459 | -34.5138 | -16.2468
9.5703 | -34.4744 | -15.2494
9.5947 | -34.5161 | -16.2968
9.6191 | -34.5124 | -15.4824
9.6436 | -34.5487 | -16.2174
9.668 | -34.5326 | -15.3683
9.6924 | -34.4906 | -16.2535
9.7168 | -34.5085 | -15.6755
9.7412 | -34.4969 | -16.2302
9.7656 | -34.5659 | -15.5077
9.79 | -34.4597 | -16.3378
9.8145 | -34.587 | -15.633
9.8389 | -34.4681 | -16.5965
9.8633 | -34.5969 | -15.9757
9.8877 | -34.4472 | -16.1604
9.9121 | -34.5995 | -16.1552
9.9365 | -34.4754 | -16.1815
9.9609 | -34.6163 | -16.2132
9.9854 | -34.5021 | -16.0798
10.0098 | -34.6188 | -16.793

## SEALED_NOISE
dBPa/√Hz | Frequency
0.00964 | -32.6539
0.009961 | -33.5991
0.010283 | -34.035
0.010604 | -35.2435
0.010925 | -35.868
0.011246 | -35.9614
0.011568 | -35.7278
0.011889 | -34.835
0.01221 | -33.9116
0.012532 | -33.6728
0.012853 | -34.9128
0.013174 | -36.6121
0.013496 | -37.3045
0.013817 | -38.0711
0.014138 | -36.9781
0.01446 | -36.1341
0.014781 | -37.4456
0.015102 | -39.01
0.015424 | -36.9046
0.015745 | -35.049
0.016066 | -34.3987
0.016388 | -34.4844
0.016709 | -34.1107
0.01703 | -34.6501
0.017352 | -36.7281
0.017673 | -37.7528
0.017994 | -36.0065
0.018316 | -34.8553
0.018637 | -34.978
0.018958 | -35.8528
0.01928 | -36.0641
0.019601 | -35.5221
0.019922 | -35.4554
0.020244 | -35.8942
0.020565 | -36.2232
0.020886 | -36.7971
0.021208 | -38.2024
0.021529 | -39.9135
0.02185 | -40.9954
0.022172 | -41.2581
0.022493 | -40.4429
0.022814 | -39.282
0.023136 | -38.8732
0.023457 | -39.2783
0.023778 | -39.9712
0.0241 | -40.8992
0.024421 | -41.4367
0.024742 | -40.1827
0.025064 | -38.5864
0.025385 | -37.8112
0.025706 | -37.5745
0.026028 | -37.949
0.026349 | -39.3466
0.02667 | -41.594
0.026992 | -43.7389
0.027313 | -44.6403
0.027634 | -43.4212
0.027956 | -42.256
0.028277 | -41.0825
0.028598 | -40.2416
0.02892 | -40.0669
0.029241 | -40.6601
0.029562 | -41.8739
0.029884 | -42.9601
0.030205 | -42.9898
0.030526 | -42.3003
0.030848 | -41.9865
0.031169 | -42.3323
0.03149 | -43.1259
0.031812 | -43.7746
0.032133 | -44.074
0.032454 | -44.2152
0.032776 | -44.1159
0.033097 | -43.4197
0.033418 | -42.2873
0.033739 | -41.5837
0.034061 | -41.8445
0.034382 | -43.1448
0.034703 | -45.1131
0.035025 | -46.6333
0.035346 | -46.7426
0.035667 | -46.1306
0.035989 | -45.6578
0.03631 | -45.4994
0.036631 | -45.5588
0.036953 | -45.6386
0.037274 | -45.7418
0.037595 | -45.8731
0.037917 | -46.0764
0.038238 | -46.3888
0.038559 | -46.7164
0.038881 | -46.8874
0.039202 | -46.7356
0.039523 | -46.2805
0.039845 | -45.7455
0.040166 | -45.4318
0.040487 | -45.3281
0.040809 | -45.4685
0.04113 | -45.9009
0.041451 | -46.6756
0.041773 | -47.7609
0.042094 | -49.0075
0.042415 | -49.765
0.042737 | -49.5721
0.043058 | -48.7176
0.043379 | -47.6822
0.043701 | -46.7066
0.044022 | -45.8087
0.044343 | -45.1094
0.044665 | -44.7284
0.044986 | -44.7715
0.045307 | -45.2695
0.045629 | -46.1682
0.04595 | -47.4209
0.046271 | -48.4082
0.046593 | -49.5364
0.046914 | -50.2463
0.047235 | -50.2939
0.047557 | -49.8965
0.047878 | -49.5767
0.048199 | -49.4853
0.048521 | -49.7863
0.048842 | -50.4575
0.049163 | -51.3223
0.049485 | -52.2097
0.049806 | -52.757
0.050127 | -52.8572
0.050449 | -52.4837
0.05077 | -51.9686
0.051091 | -51.4994
0.051413 | -51.1817
0.051734 | -51.0574
0.052055 | -51.0098
0.052377 | -51.0328
0.052698 | -51.075
0.053019 | -51.1833
0.053341 | -51.4574
0.053662 | -51.7985
0.053983 | -52.2005
0.054305 | -52.576
0.054626 | -52.8833
0.054947 | -53.2797
0.055269 | -53.8799
0.05559 | -54.478
0.055911 | -54.7325
0.056232 | -54.5404
0.056554 | -53.9787
0.056875 | -53.0438
0.057196 | -52.0246
0.057518 | -51.0598
0.057839 | -50.3298
0.05816 | -49.889
0.058482 | -49.7425
0.058803 | -49.8899
0.059124 | -50.2412
0.059446 | -50.6798
0.059767 | -51.1123
0.060088 | -51.4731
0.06041 | -51.8423
0.060731 | -52.3335
0.061052 | -52.9219
0.061374 | -53.5821
0.061695 | -54.3222
0.062016 | -55.0983
0.062338 | -55.8287
0.062659 | -56.3766
0.06298 | -56.7881
0.063302 | -57.0187
0.063623 | -56.9612
0.063944 | -56.9356
0.064266 | -56.9764
0.064587 | -57.0543
0.064908 | -57.0975
0.06523 | -57.0722
0.065551 | -56.8932
0.065872 | -56.626
0.066194 | -56.279
0.066515 | -55.9306
0.066836 | -55.6323
0.067158 | -55.4198
0.067479 | -55.3376
0.0678 | -55.3733
0.068122 | -55.4977
0.068443 | -55.6958
0.068764 | -55.9644
0.069086 | -56.2518
0.069407 | -56.6135
0.069728 | -57.0349
0.07005 | -57.4373
0.070371 | -57.85
0.070692 | -58.2395
0.071014 | -58.4007
0.071335 | -58.4205
0.071656 | -58.2444
0.071978 | -57.9221
0.072299 | -57.5551
0.07262 | -57.1455
0.072942 | -56.7967
0.073263 | -56.5044
0.073584 | -56.3003
0.073906 | -56.1982
0.074227 | -56.0792
0.074548 | -55.984
0.07487 | -55.8675
0.075191 | -55.7304
0.075512 | -55.5785
0.075834 | -55.3943
0.076155 | -55.1813
0.076476 | -54.9925
0.076798 | -54.8998
0.077119 | -54.8922
0.07744 | -55.033
0.077761 | -55.3268
0.078083 | -55.7534
0.078404 | -56.267
0.078725 | -56.8257
0.079047 | -57.3536
0.079368 | -57.7682
0.079689 | -58.0252
0.080011 | -58.0256
0.080332 | -57.7759
0.080653 | -57.457
0.080975 | -57.175
0.081296 | -56.9522
0.081617 | -56.8686
0.081939 | -56.9403
0.08226 | -57.1585
0.082581 | -57.5189
0.082903 | -57.962
0.083224 | -58.3172
0.083545 | -58.6552
0.083867 | -58.9718
0.084188 | -59.0195
0.084509 | -58.9725
0.084831 | -58.859
0.085152 | -58.5959
0.085473 | -58.2457
0.085795 | -57.854
0.086116 | -57.4873
0.086437 | -57.2109
0.086759 | -57.0063
0.08708 | -56.8956
0.087401 | -56.8901
0.087723 | -56.9573
0.088044 | -57.0775
0.088365 | -57.2066
0.088687 | -57.3143
0.089008 | -57.3417
0.089329 | -57.2538
0.089651 | -57.115
0.089972 | -56.9556
0.090293 | -56.7822
0.090615 | -56.6396
0.090936 | -56.6325
0.091257 | -56.6494
0.091579 | -56.7337
0.0919 | -56.9044
0.092221 | -57.1628
0.092543 | -57.5657
0.092864 | -57.933
0.093185 | -58.3709
0.093507 | -58.8637
0.093828 | -59.4326
0.094149 | -60.0621
0.094471 | -60.7012
0.094792 | -61.3853
0.095113 | -62.0776
0.095435 | -62.7632
0.095756 | -63.3282
0.096077 | -63.7748
0.096399 | -64.1822
0.09672 | -64.3668
0.097041 | -64.3805
0.097363 | -64.3395
0.097684 | -64.1696
0.098005 | -63.9559
0.098327 | -63.7346
0.098648 | -63.4986
0.098969 | -63.3358
0.099291 | -63.1916
0.099612 | -63.0895
0.099933 | -63.0432
0.10025 | -63.0822
0.10058 | -63.1837
0.1009 | -63.3208
0.10122 | -63.5264
0.10154 | -63.7153
0.10186 | -63.8618
0.10218 | -63.9966
0.1025 | -64.0218
0.10282 | -64.0272
0.10315 | -63.9767
0.10347 | -63.8749
0.10379 | -63.73
0.10411 | -63.5669
0.10443 | -63.3393
0.10475 | -63.1553
0.10507 | -63.0309
0.1054 | -62.8631
0.10572 | -62.5905
0.10604 | -62.4337
0.10636 | -62.2839
0.10668 | -62.1107
0.107 | -61.8849
0.10732 | -61.654
0.10765 | -61.4056
0.10797 | -61.1371
0.10829 | -60.8722
0.10861 | -60.6435
0.10893 | -60.4678
0.10925 | -60.3249
0.10957 | -60.2226
0.10989 | -60.195
0.11022 | -60.2376
0.11054 | -60.3328
0.11086 | -60.4229
0.11118 | -60.5492
0.1115 | -60.6565
0.11182 | -60.7697
0.11214 | -60.8892
0.11246 | -61.0049
0.11279 | -61.1617
0.11311 | -61.3385
0.11343 | -61.5432
0.11375 | -61.8252
0.11407 | -62.1818
0.11439 | -62.4444
0.11471 | -62.728
0.11504 | -63.0386
0.11536 | -63.3544
0.11568 | -63.6348
0.116 | -63.8705
0.11632 | -64.0674
0.11664 | -64.2591
0.11696 | -64.3575
0.11729 | -64.2634
0.11761 | -64.1237
0.11793 | -63.9671
0.11825 | -63.7033
0.11857 | -63.3892
0.11889 | -63.0459
0.11921 | -62.7144
0.11953 | -62.3817
0.11986 | -62.0804
0.12018 | -61.8247
0.1205 | -61.7743
0.12082 | -61.6762
0.12114 | -61.6258
0.12146 | -61.6436
0.12178 | -61.6861
0.12211 | -61.7733
0.12243 | -61.8991
0.12275 | -62.057
0.12307 | -62.1991
0.12339 | -62.3364
0.12371 | -62.4252
0.12403 | -62.4451
0.12435 | -62.4198
0.12467 | -62.3305
0.125 | -62.1984
0.12532 | -62.0486
0.12564 | -61.887
0.12596 | -61.7722
0.12628 | -61.7298
0.1266 | -61.7151
0.12693 | -61.7197
0.12725 | -61.7791
0.12757 | -61.8812
0.12789 | -62.0294
0.12821 | -62.2362
0.12853 | -62.4593
0.12885 | -62.6906
0.12917 | -62.9159
0.12949 | -63.0469
0.12982 | -63.2147
0.13014 | -63.3483
0.13046 | -63.4703
0.13078 | -63.5851
0.1311 | -63.6673
0.13142 | -63.7279
0.13175 | -63.7271
0.13207 | -63.7125
0.13239 | -63.6904
0.13271 | -63.6942
0.13303 | -63.6888
0.13335 | -63.6595
0.13367 | -63.6289
0.13399 | -63.6413
0.13431 | -63.6626
0.13464 | -63.7048
0.13496 | -63.7512
0.13528 | -63.7879
0.1356 | -63.8151
0.13592 | -63.8506
0.13624 | -63.8827
0.13656 | -63.9153
0.13689 | -63.9665
0.13721 | -64.0623
0.13753 | -64.1675
0.13785 | -64.2475
0.13817 | -64.3367
0.13849 | -64.4266
0.13881 | -64.5535
0.13914 | -64.6677
0.13946 | -64.7868
0.13978 | -64.8907
0.1401 | -64.9944
0.14042 | -65.0858
0.14074 | -65.1329
0.14106 | -65.1691
0.14139 | -65.1854
0.14171 | -65.2107
0.14203 | -65.2636
0.14235 | -65.2889
0.14267 | -65.3132
0.14299 | -65.361
0.14331 | -65.4256
0.14363 | -65.5577
0.14395 | -65.7061
0.14428 | -65.8116
0.1446 | -65.9373
0.14492 | -66.0907
0.14524 | -66.2656
0.14556 | -66.4466
0.14588 | -66.585
0.1462 | -66.6653
0.14653 | -66.741
0.14685 | -66.7811
0.14717 | -66.7675
0.14749 | -66.7303
0.14781 | -66.6686
0.14813 | -66.4045
0.14845 | -66.2986
0.14877 | -66.2034
0.1491 | -66.1003
0.14942 | -65.9426
0.14974 | -65.799
0.15006 | -65.6522
0.15038 | -65.5114
0.1507 | -65.3855
0.15102 | -65.257
0.15135 | -65.1538
0.15167 | -65.076
0.15199 | -65.0106
0.15231 | -64.9623
0.15263 | -64.9019
0.15295 | -64.8433
0.15327 | -64.8156
0.1536 | -64.8071
0.15392 | -64.8051
0.15424 | -64.8026
0.15456 | -64.8118
0.15488 | -64.8337
0.1552 | -64.8512
0.15552 | -64.8685
0.15584 | -64.8804
0.15617 | -64.8896
0.15649 | -64.9062
0.15681 | -64.9422
0.15713 | -64.9775
0.15745 | -65.0467
0.15777 | -65.102
0.15809 | -65.1732
0.15842 | -65.2567
0.15874 | -65.3481
0.15906 | -65.4505
0.15938 | -65.5998
0.1597 | -65.7573
0.16002 | -65.8737
0.16034 | -66.0021
0.16066 | -66.1213
0.16099 | -66.2242
0.16131 | -66.3345
0.16163 | -66.4501
0.16195 | -66.5369
0.16227 | -66.6199
0.16259 | -66.7202
0.16291 | -66.7789
0.16323 | -66.8069
0.16356 | -66.8601
0.16388 | -66.9236
0.1642 | -67.011
0.16452 | -67.1079
0.16484 | -67.1803
0.16516 | -67.2576
0.16548 | -67.351
0.16581 | -67.4827
0.16613 | -67.6093
0.16645 | -67.7451
0.16677 | -67.9298
0.16709 | -68.0972
0.16741 | -68.2467
0.16773 | -68.4108
0.16806 | -68.5974
0.16838 | -68.7983
0.1687 | -69.0027
0.16902 | -69.2209
0.16934 | -69.4409
0.16966 | -69.6345
0.16998 | -69.8059
0.1703 | -69.9355
0.17062 | -70.0388
0.17095 | -70.1599
0.17127 | -70.277
0.17159 | -70.3464
0.17191 | -70.3727
0.17223 | -70.3693
0.17255 | -70.3368
0.17288 | -70.309
0.1732 | -70.2795
0.17352 | -70.1801
0.17384 | -70.0785
0.17416 | -69.9931
0.17448 | -69.8927
0.1748 | -69.7938
0.17512 | -69.6987
0.17544 | -69.6223
0.17577 | -69.5304
0.17609 | -69.473
0.17641 | -69.4082
0.17673 | -69.316
0.17705 | -69.2246
0.17737 | -69.1326
0.17769 | -69.0596
0.17802 | -68.9907
0.17834 | -68.9134
0.17866 | -68.8287
0.17898 | -68.7516
0.1793 | -68.6572
0.17962 | -68.5605
0.17994 | -68.4602
0.18027 | -68.3625
0.18059 | -68.2433
0.18091 | -68.1167
0.18123 | -67.9937
0.18155 | -67.8906
0.18187 | -67.7931
0.18219 | -67.6911
0.18252 | -67.6057
0.18284 | -67.546
0.18316 | -67.5098
0.18348 | -67.4904
0.1838 | -67.4565
0.18412 | -67.433
0.18444 | -67.4338
0.18476 | -67.4421
0.18508 | -67.4667
0.18541 | -67.5413
0.18573 | -67.5888
0.18605 | -67.6408
0.18637 | -67.6996
0.18669 | -67.7705
0.18701 | -67.8298
0.18734 | -67.9071
0.18766 | -67.9793
0.18798 | -68.0576
0.1883 | -68.1558
0.18862 | -68.2612
0.18894 | -68.3688
0.18926 | -68.4885
0.18958 | -68.6091
0.1899 | -68.7114
0.19023 | -68.8481
0.19055 | -68.9849
0.19087 | -69.127
0.19119 | -69.2345
0.19151 | -69.3396
0.19183 | -69.4502
0.19215 | -69.5578
0.19248 | -69.6449
0.1928 | -69.74
0.19312 | -69.8345
0.19344 | -69.9022
0.19376 | -69.9792
0.19408 | -70.0664
0.1944 | -70.1251
0.19473 | -70.2066
0.19505 | -70.2898
0.19537 | -70.3564
0.19569 | -70.4194
0.19601 | -70.5123
0.19633 | -70.6029
0.19665 | -70.6712
0.19697 | -70.7653
0.1973 | -70.855
0.19762 | -70.9514
0.19794 | -71.0651
0.19826 | -71.1594
0.19858 | -71.2741
0.1989 | -71.3682
0.19922 | -71.4412
0.19955 | -71.5121
0.19987 | -71.5659
0.20019 | -71.6123
0.20051 | -71.6506
0.20083 | -71.6698
0.20115 | -71.7093
0.20147 | -71.6806
0.20179 | -71.6241
0.20212 | -71.5624
0.20244 | -71.4836
0.20276 | -71.377
0.20308 | -71.2608
0.2034 | -71.1229
0.20372 | -70.9504
0.20404 | -70.8284
0.20436 | -70.72
0.20469 | -70.6054
0.20501 | -70.4612
0.20533 | -70.3292
0.20565 | -70.2311
0.20597 | -70.1458
0.20629 | -70.0605
0.20661 | -69.9897
0.20694 | -69.9273
0.20726 | -69.8908
0.20758 | -69.8344
0.2079 | -69.7852
0.20822 | -69.7578
0.20854 | -69.7488
0.20886 | -69.7562
0.20919 | -69.7719
0.20951 | -69.7824
0.20983 | -69.8006
0.21015 | -69.8237
0.21047 | -69.8548
0.21079 | -69.8593
0.21111 | -69.8545
0.21143 | -69.8503
0.21175 | -69.8264
0.21208 | -69.7808
0.2124 | -69.7345
0.21272 | -69.6898
0.21304 | -69.595
0.21336 | -69.5448
0.21368 | -69.4991
0.21401 | -69.4581
0.21433 | -69.4133
0.21465 | -69.3775
0.21497 | -69.3542
0.21529 | -69.3375
0.21561 | -69.3287
0.21593 | -69.3203
0.21625 | -69.318
0.21657 | -69.3358
0.2169 | -69.374
0.21722 | -69.4102
0.21754 | -69.4541
0.21786 | -69.5172
0.21818 | -69.5758
0.2185 | -69.6324
0.21882 | -69.711
0.21915 | -69.7944
0.21947 | -69.8893
0.21979 | -69.9857
0.22011 | -70.0817
0.22043 | -70.1967
0.22075 | -70.3039
0.22107 | -70.4007
0.2214 | -70.4907
0.22172 | -70.5652
0.22204 | -70.649
0.22236 | -70.6815
0.22268 | -70.7626
0.223 | -70.8264
0.22332 | -70.9142
0.22365 | -70.9893
0.22397 | -71.0286
0.22429 | -71.0424
0.22461 | -71.0479
0.22493 | -71.0479
0.22525 | -71.0576
0.22557 | -71.0559
0.22589 | -71.0514
0.22621 | -71.0519
0.22654 | -71.0708
0.22686 | -71.0799
0.22718 | -71.091
0.2275 | -71.1203
0.22782 | -71.1331
0.22814 | -71.11
0.22847 | -71.1135
0.22879 | -71.146
0.22911 | -71.1687
0.22943 | -71.1701
0.22975 | -71.1745
0.23007 | -71.2027
0.23039 | -71.2161
0.23071 | -71.216
0.23103 | -71.2057
0.23136 | -71.2125
0.23168 | -71.1923
0.232 | -71.1718
0.23232 | -71.1612
0.23264 | -71.1422
0.23296 | -71.1138
0.23328 | -71.0888
0.23361 | -71.048
0.23393 | -71.0043
0.23425 | -70.9598
0.23457 | -70.9281
0.23489 | -70.9014
0.23521 | -70.854
0.23553 | -70.8039
0.23586 | -70.7605
0.23618 | -70.7279
0.2365 | -70.7078
0.23682 | -70.6888
0.23714 | -70.6713
0.23746 | -70.6584
0.23778 | -70.6623
0.2381 | -70.6898
0.23843 | -70.711
0.23875 | -70.7447
0.23907 | -70.7867
0.23939 | -70.8302
0.23971 | -70.8892
0.24003 | -70.9573
0.24035 | -71.042
0.24068 | -71.1254
0.241 | -71.2197
0.24132 | -71.2932
0.24164 | -71.3771
0.24196 | -71.4602
0.24228 | -71.5453
0.2426 | -71.6283
0.24292 | -71.711
0.24325 | -71.7872
0.24357 | -71.8701
0.24389 | -71.9493
0.24421 | -72.024
0.24453 | -72.095
0.24485 | -72.1759
0.24517 | -72.277
0.24549 | -72.3372
0.24582 | -72.3934
0.24614 | -72.4483
0.24646 | -72.5064
0.24678 | -72.5558
0.2471 | -72.5883
0.24742 | -72.6488
0.24774 | -72.7014
0.24807 | -72.7574
0.24839 | -72.7895
0.24871 | -72.82
0.24903 | -72.8423
0.24935 | -72.8854
0.24967 | -72.9285
0.24999 | -72.9487
0.25032 | -72.92
0.25064 | -72.9212
0.25096 | -72.8908
0.25128 | -72.8706
0.2516 | -72.8575
0.25192 | -72.8346
0.25224 | -72.7931
0.25256 | -72.7558
0.25289 | -72.7091
0.25321 | -72.6614
0.25353 | -72.6066
0.25385 | -72.5325
0.25417 | -72.4468
0.25449 | -72.3809
0.25481 | -72.3206
0.25514 | -72.2631
0.25546 | -72.2007
0.25578 | -72.13
0.2561 | -72.0638
0.25642 | -71.9934
0.25674 | -71.9357
0.25706 | -71.8844
0.25738 | -71.8204
0.25771 | -71.7656
0.25803 | -71.7188
0.25835 | -71.6841
0.25867 | -71.6574
0.25899 | -71.6446
0.25931 | -71.6792
0.25963 | -71.6888
0.25995 | -71.6958
0.26028 | -71.7035
0.2606 | -71.7233
0.26092 | -71.7571
0.26124 | -71.7971
0.26156 | -71.8507
0.26188 | -71.9032
0.2622 | -71.9624
0.26253 | -72.036
0.26285 | -72.1034
0.26317 | -72.1844
0.26349 | -72.2685
0.26381 | -72.3472
0.26413 | -72.4264
0.26445 | -72.5239
0.26477 | -72.6158
0.2651 | -72.7118
0.26542 | -72.8072
0.26574 | -72.9002
0.26606 | -72.9943
0.26638 | -73.0828
0.2667 | -73.1658
0.26702 | -73.2562
0.26734 | -73.3618
0.26767 | -73.4458
0.26799 | -73.5125
0.26831 | -73.5802
0.26863 | -73.6322
0.26895 | -73.7117
0.26927 | -73.7913
0.26959 | -73.866
0.26992 | -73.9561
0.27024 | -74.026
0.27056 | -74.076
0.27088 | -74.1274
0.2712 | -74.195
0.27152 | -74.2612
0.27184 | -74.3105
0.27216 | -74.347
0.27249 | -74.374
0.27281 | -74.4326
0.27313 | -74.4992
0.27345 | -74.5301
0.27377 | -74.5623
0.27409 | -74.59
0.27442 | -74.6198
0.27474 | -74.6558
0.27506 | -74.6866
0.27538 | -74.7244
0.2757 | -74.7615
0.27602 | -74.79
0.27634 | -74.8299
0.27666 | -74.8688
0.27698 | -74.9007
0.27731 | -74.9219
0.27763 | -74.9407
0.27795 | -74.983
0.27827 | -75.017
0.27859 | -75.0585
0.27891 | -75.0952
0.27923 | -75.1243
0.27956 | -75.1639
0.27988 | -75.2175
0.2802 | -75.2357
0.28052 | -75.2455
0.28084 | -75.266
0.28116 | -75.2647
0.28148 | -75.2686
0.28181 | -75.2824
0.28213 | -75.2973
0.28245 | -75.3185
0.28277 | -75.3074
0.28309 | -75.3058
0.28341 | -75.2703
0.28373 | -75.2521
0.28405 | -75.215
0.28438 | -75.1657
0.2847 | -75.1194
0.28502 | -75.058
0.28534 | -74.9941
0.28566 | -74.915
0.28598 | -74.8288
0.2863 | -74.7416
0.28663 | -74.6475
0.28695 | -74.5469
0.28727 | -74.4409
0.28759 | -74.3568
0.28791 | -74.2804
0.28823 | -74.2093
0.28855 | -74.1365
0.28887 | -74.063
0.2892 | -74.003
0.28952 | -73.9422
0.28984 | -73.883
0.29016 | -73.8315
0.29048 | -73.7772
0.2908 | -73.732
0.29112 | -73.6994
0.29145 | -73.6548
0.29177 | -73.6291
0.29209 | -73.619
0.29241 | -73.6262
0.29273 | -73.6419
0.29305 | -73.6575
0.29337 | -73.6731
0.29369 | -73.6968
0.29402 | -73.7214
0.29434 | -73.7498
0.29466 | -73.7889
0.29498 | -73.8247
0.2953 | -73.8703
0.29562 | -73.9105
0.29594 | -73.9567
0.29627 | -73.9865
0.29659 | -74.008
0.29691 | -74.0394
0.29723 | -74.0713
0.29755 | -74.1087
0.29787 | -74.1433
0.29819 | -74.1601
0.29851 | -74.1563
0.29884 | -74.1559
0.29916 | -74.1649
0.29948 | -74.1704
0.2998 | -74.1683
0.30012 | -74.1699
0.30044 | -74.1684
0.30076 | -74.1588
0.30108 | -74.1538
0.30141 | -74.1494
0.30173 | -74.1382
0.30205 | -74.1366
0.30237 | -74.1241
0.30269 | -74.0912
0.30301 | -74.0526
0.30333 | -74.0292
0.30366 | -74.0091
0.30398 | -73.9769
0.3043 | -73.9674
0.30462 | -73.9413
0.30494 | -73.9204
0.30526 | -73.9103
0.30558 | -73.8875
0.3059 | -73.8989
0.30623 | -73.8968
0.30655 | -73.891
0.30687 | -73.8916
0.30719 | -73.8942
0.30751 | -73.8832
0.30783 | -73.8463
0.30815 | -73.8334
0.30847 | -73.8336
0.3088 | -73.8288
0.30912 | -73.8151
0.30944 | -73.8102
0.30976 | -73.8111
0.31008 | -73.8063
0.3104 | -73.799
0.31072 | -73.7893
0.31105 | -73.7878
0.31137 | -73.7864
0.31169 | -73.7884
0.31201 | -73.7997
0.31233 | -73.8034
0.31265 | -73.7895
0.31297 | -73.7901
0.31329 | -73.7867
0.31362 | -73.7776
0.31394 | -73.763
0.31426 | -73.7539
0.31458 | -73.7435
0.3149 | -73.7479
0.31522 | -73.7466
0.31555 | -73.7456
0.31587 | -73.7417
0.31619 | -73.7479
0.31651 | -73.7481
0.31683 | -73.7576
0.31715 | -73.766
0.31747 | -73.7792
0.31779 | -73.7897
0.31811 | -73.8199
0.31844 | -73.8545
0.31876 | -73.8816
0.31908 | -73.9215
0.3194 | -73.9641
0.31972 | -73.9992
0.32004 | -74.0228
0.32037 | -74.0626
0.32069 | -74.118
0.32101 | -74.1773
0.32133 | -74.222
0.32165 | -74.2701
0.32197 | -74.3316
0.32229 | -74.3951
0.32261 | -74.4572
0.32294 | -74.5027
0.32326 | -74.5511
0.32358 | -74.6066
0.3239 | -74.6555
0.32422 | -74.6894
0.32454 | -74.7325
0.32486 | -74.7751
0.32518 | -74.8279
0.32551 | -74.8749
0.32583 | -74.9099
0.32615 | -74.9529
0.32647 | -74.9761
0.32679 | -74.9939
0.32711 | -75.0141
0.32743 | -75.021
0.32776 | -75.029
0.32808 | -75.044
0.3284 | -75.0467
0.32872 | -75.0417
0.32904 | -75.0277
0.32936 | -75.0083
0.32968 | -75.0052
0.33 | -75.0077
0.33033 | -75.0174
0.33065 | -75.0422
0.33097 | -75.0305
0.33129 | -75.0014
0.33161 | -74.9876
0.33193 | -74.9737
0.33225 | -74.9478
0.33258 | -74.9338
0.3329 | -74.9138
0.33322 | -74.8984
0.33354 | -74.8991
0.33386 | -74.8968
0.33418 | -74.8811
0.3345 | -74.8668
0.33482 | -74.8526
0.33515 | -74.8378
0.33547 | -74.8326
0.33579 | -74.8368
0.33611 | -74.8424
0.33643 | -74.8385
0.33675 | -74.8322
0.33707 | -74.8192
0.3374 | -74.807
0.33772 | -74.7885
0.33804 | -74.7656
0.33836 | -74.7525
0.33868 | -74.7409
0.339 | -74.7262
0.33932 | -74.7187
0.33964 | -74.7069
0.33997 | -74.687
0.34029 | -74.6724
0.34061 | -74.6256
0.34093 | -74.592
0.34125 | -74.5741
0.34157 | -74.5588
0.34189 | -74.5425
0.34221 | -74.5241
0.34254 | -74.5105
0.34286 | -74.5102
0.34318 | -74.5029
0.3435 | -74.5001
0.34382 | -74.5152
0.34414 | -74.5177
0.34446 | -74.5211
0.34479 | -74.5181
0.34511 | -74.5175
0.34543 | -74.4898
0.34575 | -74.4723
0.34607 | -74.4685
0.34639 | -74.4675
0.34671 | -74.4872
0.34703 | -74.5089
0.34736 | -74.5081
0.34768 | -74.499
0.348 | -74.4923
0.34832 | -74.493
0.34864 | -74.4902
0.34896 | -74.4946
0.34928 | -74.4914
0.3496 | -74.4893
0.34993 | -74.4858
0.35025 | -74.4783
0.35057 | -74.4729
0.35089 | -74.4634
0.35121 | -74.4549
0.35153 | -74.4534
0.35185 | -74.447
0.35218 | -74.4419
0.3525 | -74.4351
0.35282 | -74.4337
0.35314 | -74.4331
0.35346 | -74.4427
0.35378 | -74.4566
0.3541 | -74.4592
0.35442 | -74.46
0.35475 | -74.4585
0.35507 | -74.4625
0.35539 | -74.4746
0.35571 | -74.4818
0.35603 | -74.4807
0.35635 | -74.487
0.35668 | -74.4987
0.357 | -74.5024
0.35732 | -74.5209
0.35764 | -74.5413
0.35796 | -74.5516
0.35828 | -74.567
0.3586 | -74.5825
0.35892 | -74.6045
0.35924 | -74.6284
0.35957 | -74.6437
0.35989 | -74.6603
0.36021 | -74.6808
0.36053 | -74.7099
0.36085 | -74.735
0.36117 | -74.7634
0.3615 | -74.7918
0.36182 | -74.8236
0.36214 | -74.8582
0.36246 | -74.8925
0.36278 | -74.9363
0.3631 | -74.9851
0.36342 | -75.0228
0.36374 | -75.0611
0.36407 | -75.1046
0.36439 | -75.1526
0.36471 | -75.1923
0.36503 | -75.2335
0.36535 | -75.2738
0.36567 | -75.3496
0.36599 | -75.4112
0.36631 | -75.4525
0.36664 | -75.4874
0.36696 | -75.5152
0.36728 | -75.5498
0.3676 | -75.5887
0.36792 | -75.6186
0.36824 | -75.6474
0.36856 | -75.6793
0.36889 | -75.709
0.36921 | -75.7351
0.36953 | -75.7716
0.36985 | -75.7978
0.37017 | -75.8202
0.37049 | -75.8473
0.37081 | -75.8532
0.37113 | -75.9105
0.37146 | -75.9426
0.37178 | -75.9575
0.3721 | -75.9723
0.37242 | -75.9871
0.37274 | -76.0051
0.37306 | -76.0239
0.37338 | -76.0497
0.37371 | -76.0593
0.37403 | -76.055
0.37435 | -76.0547
0.37467 | -76.0553
0.37499 | -76.0677
0.37531 | -76.082
0.37563 | -76.0944
0.37595 | -76.1043
0.37628 | -76.1055
0.3766 | -76.1147
0.37692 | -76.1297
0.37724 | -76.1364
0.37756 | -76.1428
0.37788 | -76.1441
0.3782 | -76.1538
0.37853 | -76.1612
0.37885 | -76.1631
0.37917 | -76.169
0.37949 | -76.1701
0.37981 | -76.1764
0.38013 | -76.1753
0.38045 | -76.1799
0.38077 | -76.1842
0.3811 | -76.1932
0.38142 | -76.206
0.38174 | -76.2109
0.38206 | -76.2152
0.38238 | -76.2257
0.3827 | -76.2431
0.38302 | -76.2579
0.38334 | -76.2653
0.38367 | -76.2703
0.38399 | -76.2778
0.38431 | -76.289
0.38463 | -76.2979
0.38495 | -76.3082
0.38527 | -76.3268
0.38559 | -76.3448
0.38592 | -76.3709
0.38624 | -76.4086
0.38656 | -76.4346
0.38688 | -76.4609
0.3872 | -76.4972
0.38752 | -76.532
0.38784 | -76.5892
0.38816 | -76.6311
0.38849 | -76.6719
0.38881 | -76.7043
0.38913 | -76.7425
0.38945 | -76.7821
0.38977 | -76.8248
0.39009 | -76.8709
0.39041 | -76.9194
0.39073 | -76.9694
0.39106 | -77.0206
0.39138 | -77.0628
0.3917 | -77.0959
0.39202 | -77.1465
0.39234 | -77.1966
0.39266 | -77.2433
0.39298 | -77.2768
0.39331 | -77.3037
0.39363 | -77.3355
0.39395 | -77.3807
0.39427 | -77.4113
0.39459 | -77.4343
0.39491 | -77.4695
0.39523 | -77.5054
0.39555 | -77.5235
0.39588 | -77.5402
0.3962 | -77.5521
0.39652 | -77.5604
0.39684 | -77.5693
0.39716 | -77.5762
0.39748 | -77.5798
0.39781 | -77.5849
0.39813 | -77.5981
0.39845 | -77.5907
0.39877 | -77.5892
0.39909 | -77.579
0.39941 | -77.5602
0.39973 | -77.5478
0.40005 | -77.531
0.40037 | -77.5083
0.4007 | -77.4847
0.40102 | -77.4733
0.40134 | -77.4761
0.40166 | -77.4859
0.40198 | -77.493
0.4023 | -77.4811
0.40263 | -77.4707
0.40295 | -77.4558
0.40327 | -77.4499
0.40359 | -77.4519
0.40391 | -77.4662
0.40423 | -77.4677
0.40455 | -77.4591
0.40487 | -77.4435
0.4052 | -77.425
0.40552 | -77.4244
0.40584 | -77.4071
0.40616 | -77.38
0.40648 | -77.3705
0.4068 | -77.3658
0.40712 | -77.363
0.40745 | -77.3537
0.40777 | -77.3573
0.40809 | -77.3484
0.40841 | -77.3384
0.40873 | -77.3315
0.40905 | -77.3248
0.40937 | -77.3169
0.40969 | -77.3135
0.41002 | -77.2902
0.41034 | -77.2665
0.41066 | -77.2431
0.41098 | -77.2258
0.4113 | -77.2056
0.41162 | -77.178
0.41194 | -77.1559
0.41226 | -77.1368
0.41259 | -77.1185
0.41291 | -77.0987
0.41323 | -77.07
0.41355 | -77.0327
0.41387 | -76.9982
0.41419 | -76.9685
0.41451 | -76.9196
0.41484 | -76.8736
0.41516 | -76.8291
0.41548 | -76.7882
0.4158 | -76.7302
0.41612 | -76.6676
0.41644 | -76.616
0.41676 | -76.5629
0.41708 | -76.5054
0.41741 | -76.4712
0.41773 | -76.4259
0.41805 | -76.3638
0.41837 | -76.3016
0.41869 | -76.2499
0.41901 | -76.1902
0.41933 | -76.1243
0.41966 | -76.068
0.41998 | -76.0176
0.4203 | -75.9699
0.42062 | -75.9301
0.42094 | -75.8829
0.42126 | -75.827
0.42158 | -75.7654
0.4219 | -75.7098
0.42223 | -75.6619
0.42255 | -75.6126
0.42287 | -75.562
0.42319 | -75.5144
0.42351 | -75.4661
0.42383 | -75.4224
0.42415 | -75.3816
0.42447 | -75.3438
0.4248 | -75.3074
0.42512 | -75.2745
0.42544 | -75.2443
0.42576 | -75.214
0.42608 | -75.1871
0.4264 | -75.1773
0.42672 | -75.1453
0.42705 | -75.1175
0.42737 | -75.0938
0.42769 | -75.0734
0.42801 | -75.0459
0.42833 | -75.0269
0.42865 | -75.0172
0.42897 | -75.0066
0.42929 | -74.9914
0.42962 | -74.9843
0.42994 | -74.9831
0.43026 | -74.9845
0.43058 | -74.983
0.4309 | -74.9812
0.43122 | -74.9816
0.43154 | -74.975
0.43187 | -74.9838
0.43219 | -74.9857
0.43251 | -74.9898
0.43283 | -75.0005
0.43315 | -75.0132
0.43347 | -75.029
0.43379 | -75.0471
0.43411 | -75.0679
0.43444 | -75.0955
0.43476 | -75.1244
0.43508 | -75.1505
0.4354 | -75.1874
0.43572 | -75.2338
0.43604 | -75.2776
0.43636 | -75.3143
0.43668 | -75.3507
0.43701 | -75.3873
0.43733 | -75.43
0.43765 | -75.467
0.43797 | -75.5039
0.43829 | -75.553
0.43861 | -75.6034
0.43894 | -75.6471
0.43926 | -75.6941
0.43958 | -75.7483
0.4399 | -75.7951
0.44022 | -75.8548
0.44054 | -75.9217
0.44086 | -75.9825
0.44118 | -76.0367
0.4415 | -76.0943
0.44183 | -76.1599
0.44215 | -76.2188
0.44247 | -76.2755
0.44279 | -76.3331
0.44311 | -76.3903
0.44343 | -76.4541
0.44376 | -76.5241
0.44408 | -76.5923
0.4444 | -76.6627
0.44472 | -76.7289
0.44504 | -76.8037
0.44536 | -76.8773
0.44568 | -76.9446
0.446 | -77.0098
0.44633 | -77.0904
0.44665 | -77.1659
0.44697 | -77.2304
0.44729 | -77.2983
0.44761 | -77.3634
0.44793 | -77.4225
0.44825 | -77.4781
0.44858 | -77.5511
0.4489 | -77.6284
0.44922 | -77.6872
0.44954 | -77.7294
0.44986 | -77.7879
0.45018 | -77.8545
0.4505 | -77.9013
0.45082 | -77.935
0.45115 | -77.9771
0.45147 | -78.0231
0.45179 | -78.0617
0.45211 | -78.101
0.45243 | -78.1509
0.45275 | -78.1975
0.45307 | -78.2361
0.45339 | -78.2628
0.45372 | -78.2793
0.45404 | -78.2841
0.45436 | -78.3083
0.45468 | -78.3238
0.455 | -78.3268
0.45532 | -78.3323
0.45564 | -78.3364
0.45597 | -78.3462
0.45629 | -78.3576
0.45661 | -78.3613
0.45693 | -78.365
0.45725 | -78.3692
0.45757 | -78.3694
0.45789 | -78.3802
0.45821 | -78.3991
0.45854 | -78.4096
0.45886 | -78.4155
0.45918 | -78.421
0.4595 | -78.4393
0.45982 | -78.4436
0.46014 | -78.4444
0.46046 | -78.4578
0.46079 | -78.4724
0.46111 | -78.4751
0.46143 | -78.4771
0.46175 | -78.4802
0.46207 | -78.4872
0.46239 | -78.4949
0.46271 | -78.5014
0.46303 | -78.5282
0.46336 | -78.5591
0.46368 | -78.5703
0.464 | -78.5835
0.46432 | -78.5976
0.46464 | -78.6126
0.46496 | -78.6305
0.46528 | -78.6472
0.4656 | -78.6686
0.46593 | -78.6895
0.46625 | -78.7165
0.46657 | -78.7451
0.46689 | -78.7715
0.46721 | -78.8013
0.46753 | -78.8304
0.46785 | -78.8561
0.46818 | -78.8892
0.4685 | -78.9212
0.46882 | -78.9402
0.46914 | -78.9512
0.46946 | -78.9703
0.46978 | -79.0023
0.4701 | -79.0302
0.47042 | -79.0512
0.47075 | -79.0806
0.47107 | -79.1138
0.47139 | -79.1387
0.47171 | -79.1575
0.47203 | -79.179
0.47235 | -79.1996
0.47267 | -79.2032
0.473 | -79.2185
0.47332 | -79.2331
0.47364 | -79.2498
0.47396 | -79.2654
0.47428 | -79.2704
0.4746 | -79.2692
0.47492 | -79.2749
0.47524 | -79.2749
0.47557 | -79.281
0.47589 | -79.2872
0.47621 | -79.277
0.47653 | -79.2607
0.47685 | -79.2473
0.47717 | -79.2345
0.47749 | -79.2226
0.47781 | -79.2069
0.47814 | -79.188
0.47846 | -79.1724
0.47878 | -79.1551
0.4791 | -79.1387
0.47942 | -79.1251
0.47974 | -79.1051
0.48007 | -79.0863
0.48039 | -79.0598
0.48071 | -79.0361
0.48103 | -79.0155
0.48135 | -78.9903
0.48167 | -78.9639
0.48199 | -78.9386
0.48231 | -78.9256
0.48263 | -78.9155
0.48296 | -78.8944
0.48328 | -78.8745
0.4836 | -78.8595
0.48392 | -78.8509
0.48424 | -78.8328
0.48456 | -78.8126
0.48489 | -78.7942
0.48521 | -78.7865
0.48553 | -78.7652
0.48585 | -78.7394
0.48617 | -78.725
0.48649 | -78.7235
0.48681 | -78.7278
0.48713 | -78.7193
0.48746 | -78.7181
0.48778 | -78.7177
0.4881 | -78.725
0.48842 | -78.732
0.48874 | -78.7336
0.48906 | -78.741
0.48938 | -78.7451
0.48971 | -78.7512
0.49003 | -78.7589
0.49035 | -78.7686
0.49067 | -78.7721
0.49099 | -78.7825
0.49131 | -78.8012
0.49163 | -78.8079
0.49195 | -78.8209
0.49228 | -78.8286
0.4926 | -78.8388
0.49292 | -78.8515
0.49324 | -78.8619
0.49356 | -78.8717
0.49388 | -78.8777
0.4942 | -78.8787
0.49452 | -78.8933
0.49485 | -78.902
0.49517 | -78.9108
0.49549 | -78.9199
0.49581 | -78.9308
0.49613 | -78.945
0.49645 | -78.9596
0.49677 | -78.9701
0.4971 | -78.9794
0.49742 | -78.9878
0.49774 | -78.9916
0.49806 | -79.0081
0.49838 | -79.0248
0.4987 | -79.0306
0.49902 | -79.0373
0.49934 | -79.0467
0.49967 | -79.0544
0.49999 | -79.0621
0.50031 | -79.0696
0.50063 | -79.0882
0.50095 | -79.1005
0.50127 | -79.1073
0.50159 | -79.1161
0.50192 | -79.1325
0.50224 | -79.1496
0.50256 | -79.1619
0.50288 | -79.1777
0.5032 | -79.207
0.50352 | -79.2436
0.50384 | -79.2697
0.50416 | -79.2876
0.50449 | -79.3054
0.50481 | -79.3281
0.50513 | -79.339
0.50545 | -79.3545
0.50577 | -79.3781
0.50609 | -79.4039
0.50641 | -79.4319
0.50674 | -79.4606
0.50706 | -79.4827
0.50738 | -79.5065
0.5077 | -79.5231
0.50802 | -79.5431
0.50834 | -79.571
0.50866 | -79.5996
0.50898 | -79.6301
0.50931 | -79.673
0.50963 | -79.7065
0.50995 | -79.7375
0.51027 | -79.7727
0.51059 | -79.806
0.51091 | -79.8377
0.51123 | -79.8732
0.51155 | -79.9133
0.51188 | -79.9395
0.5122 | -79.9698
0.51252 | -79.9998
0.51284 | -80.0237
0.51316 | -80.0481
0.51348 | -80.077
0.5138 | -80.1054
0.51413 | -80.1338
0.51445 | -80.169
0.51477 | -80.2052
0.51509 | -80.2331
0.51541 | -80.2665
0.51573 | -80.3038
0.51605 | -80.3331
0.51638 | -80.3606
0.5167 | -80.395
0.51702 | -80.4234
0.51734 | -80.446
0.51766 | -80.4629
0.51798 | -80.4781
0.5183 | -80.4954
0.51862 | -80.5129
0.51895 | -80.5224
0.51927 | -80.5388
0.51959 | -80.5553
0.51991 | -80.5761
0.52023 | -80.5986
0.52055 | -80.6228
0.52087 | -80.6369
0.52119 | -80.6607
0.52152 | -80.6879
0.52184 | -80.7019
0.52216 | -80.7083
0.52248 | -80.6991
0.5228 | -80.6938
0.52312 | -80.6982
0.52344 | -80.7013
0.52377 | -80.7077
0.52409 | -80.7162
0.52441 | -80.7221
0.52473 | -80.7205
0.52505 | -80.7206
0.52537 | -80.7232
0.52569 | -80.7284
0.52602 | -80.7354
0.52634 | -80.7478
0.52666 | -80.74
0.52698 | -80.7236
0.5273 | -80.7208
0.52762 | -80.7272
0.52794 | -80.719
0.52826 | -80.6982
0.52858 | -80.6902
0.52891 | -80.6904
0.52923 | -80.6806
0.52955 | -80.6738
0.52987 | -80.6714
0.53019 | -80.6688
0.53051 | -80.6707
0.53083 | -80.6535
0.53116 | -80.6433
0.53148 | -80.6321
0.5318 | -80.6095
0.53212 | -80.5883
0.53244 | -80.5752
0.53276 | -80.5629
0.53308 | -80.5505
0.53341 | -80.5338
0.53373 | -80.5115
0.53405 | -80.486
0.53437 | -80.4622
0.53469 | -80.4336
0.53501 | -80.4104
0.53533 | -80.3915
0.53565 | -80.3688
0.53598 | -80.3449
0.5363 | -80.3205
0.53662 | -80.3015
0.53694 | -80.2792
0.53726 | -80.2553
0.53758 | -80.2236
0.5379 | -80.1954
0.53822 | -80.1672
0.53855 | -80.1338
0.53887 | -80.1043
0.53919 | -80.0798
0.53951 | -80.052
0.53983 | -80.0246
0.54015 | -79.9938
0.54048 | -79.9567
0.5408 | -79.9306
0.54112 | -79.9062
0.54144 | -79.8857
0.54176 | -79.8546
0.54208 | -79.8319
0.5424 | -79.8052
0.54272 | -79.7724
0.54304 | -79.7406
0.54337 | -79.7151
0.54369 | -79.6854
0.54401 | -79.6604
0.54433 | -79.6401
0.54465 | -79.6214
0.54497 | -79.6001
0.54529 | -79.5703
0.54562 | -79.5449
0.54594 | -79.5263
0.54626 | -79.5091
0.54658 | -79.4916
0.5469 | -79.4846
0.54722 | -79.4717
0.54754 | -79.4517
0.54787 | -79.4348
0.54819 | -79.4226
0.54851 | -79.406
0.54883 | -79.3924
0.54915 | -79.3857
0.54947 | -79.38
0.54979 | -79.3725
0.55011 | -79.3652
0.55044 | -79.3605
0.55076 | -79.3587
0.55108 | -79.3547
0.5514 | -79.3488
0.55172 | -79.3462
0.55204 | -79.3419
0.55236 | -79.3381
0.55268 | -79.3377
0.55301 | -79.3413
0.55333 | -79.3417
0.55365 | -79.3431
0.55397 | -79.3488
0.55429 | -79.357
0.55461 | -79.3696
0.55493 | -79.3866
0.55526 | -79.4055
0.55558 | -79.417
0.5559 | -79.429
0.55622 | -79.4434
0.55654 | -79.4544
0.55686 | -79.4653
0.55718 | -79.4741
0.55751 | -79.4832
0.55783 | -79.4975
0.55815 | -79.5138
0.55847 | -79.5307
0.55879 | -79.5439
0.55911 | -79.5562
0.55943 | -79.5764
0.55975 | -79.5974
0.56008 | -79.6147
0.5604 | -79.6325
0.56072 | -79.6599
0.56104 | -79.6906
0.56136 | -79.7116
0.56168 | -79.731
0.562 | -79.7465
0.56232 | -79.7604
0.56265 | -79.7683
0.56297 | -79.7853
0.56329 | -79.805
0.56361 | -79.8108
0.56393 | -79.8256
0.56425 | -79.8293
0.56457 | -79.8163
0.5649 | -79.8261
0.56522 | -79.8149
0.56554 | -79.8232
0.56586 | -79.8265
0.56618 | -79.8286
0.5665 | -79.8309
0.56682 | -79.8391
0.56715 | -79.8425
0.56747 | -79.8391
0.56779 | -79.8335
0.56811 | -79.8327
0.56843 | -79.8336
0.56875 | -79.8298
0.56907 | -79.8262
0.56939 | -79.8177
0.56971 | -79.8114
0.57004 | -79.8042
0.57036 | -79.7868
0.57068 | -79.779
0.571 | -79.7711
0.57132 | -79.7507
0.57164 | -79.732
0.57196 | -79.7178
0.57229 | -79.7049
0.57261 | -79.6954
0.57293 | -79.6903
0.57325 | -79.6806
0.57357 | -79.6689
0.57389 | -79.6515
0.57421 | -79.6338
0.57454 | -79.619
0.57486 | -79.6042
0.57518 | -79.5905
0.5755 | -79.58
0.57582 | -79.5613
0.57614 | -79.5452
0.57646 | -79.5333
0.57678 | -79.5238
0.57711 | -79.5162
0.57743 | -79.5094
0.57775 | -79.5085
0.57807 | -79.5092
0.57839 | -79.5039
0.57871 | -79.4931
0.57903 | -79.4855
0.57935 | -79.483
0.57968 | -79.4801
0.58 | -79.4789
0.58032 | -79.4776
0.58064 | -79.4802
0.58096 | -79.4835
0.58128 | -79.4872
0.58161 | -79.4933
0.58193 | -79.502
0.58225 | -79.5159
0.58257 | -79.5285
0.58289 | -79.5336
0.58321 | -79.5473
0.58353 | -79.5635
0.58385 | -79.6066
0.58417 | -79.6267
0.5845 | -79.6522
0.58482 | -79.6795
0.58514 | -79.7124
0.58546 | -79.7425
0.58578 | -79.7694
0.5861 | -79.7967
0.58642 | -79.8264
0.58675 | -79.8602
0.58707 | -79.8943
0.58739 | -79.9328
0.58771 | -79.9755
0.58803 | -80.0185
0.58835 | -80.0598
0.58867 | -80.1077
0.589 | -80.1563
0.58932 | -80.2006
0.58964 | -80.2422
0.58996 | -80.2869
0.59028 | -80.3378
0.5906 | -80.3901
0.59092 | -80.4443
0.59124 | -80.4899
0.59157 | -80.5426
0.59189 | -80.5906
0.59221 | -80.637
0.59253 | -80.6897
0.59285 | -80.7327
0.59317 | -80.7601
0.59349 | -80.8063
0.59381 | -80.848
0.59414 | -80.8897
0.59446 | -80.9292
0.59478 | -80.97
0.5951 | -81.004
0.59542 | -81.0431
0.59574 | -81.0789
0.59606 | -81.1185
0.59639 | -81.1507
0.59671 | -81.192
0.59703 | -81.2223
0.59735 | -81.2523
0.59767 | -81.2866
0.59799 | -81.3201
0.59831 | -81.3516
0.59864 | -81.3704
0.59896 | -81.3895
0.59928 | -81.408
0.5996 | -81.4251
0.59992 | -81.4439
0.60024 | -81.4557
0.60056 | -81.4665
0.60088 | -81.4795
0.60121 | -81.4884
0.60153 | -81.4932
0.60185 | -81.4893
0.60217 | -81.4917
0.60249 | -81.4809
0.60281 | -81.4806
0.60313 | -81.4861
0.60345 | -81.4902
0.60378 | -81.5052
0.6041 | -81.5057
0.60442 | -81.5028
0.60474 | -81.5069
0.60506 | -81.502
0.60538 | -81.5049
0.6057 | -81.5304
0.60603 | -81.5333
0.60635 | -81.535
0.60667 | -81.5306
0.60699 | -81.5299
0.60731 | -81.524
0.60763 | -81.5159
0.60795 | -81.5094
0.60828 | -81.4845
0.6086 | -81.4773
0.60892 | -81.4804
0.60924 | -81.4749
0.60956 | -81.4657
0.60988 | -81.4564
0.6102 | -81.4498
0.61052 | -81.4462
0.61084 | -81.4452
0.61117 | -81.4362
0.61149 | -81.427
0.61181 | -81.4242
0.61213 | -81.4273
0.61245 | -81.4316
0.61277 | -81.4223
0.61309 | -81.4249
0.61342 | -81.4328
0.61374 | -81.4317
0.61406 | -81.4188
0.61438 | -81.4119
0.6147 | -81.416
0.61502 | -81.4061
0.61534 | -81.3974
0.61567 | -81.4019
0.61599 | -81.4037
0.61631 | -81.4099
0.61663 | -81.4139
0.61695 | -81.4171
0.61727 | -81.4197
0.61759 | -81.4254
0.61791 | -81.4319
0.61824 | -81.435
0.61856 | -81.4385
0.61888 | -81.4372
0.6192 | -81.4258
0.61952 | -81.4132
0.61984 | -81.4125
0.62016 | -81.408
0.62048 | -81.4057
0.62081 | -81.4071
0.62113 | -81.3984
0.62145 | -81.4003
0.62177 | -81.3921
0.62209 | -81.3927
0.62241 | -81.3987
0.62274 | -81.3942
0.62306 | -81.387
0.62338 | -81.3795
0.6237 | -81.3739
0.62402 | -81.3635
0.62434 | -81.3553
0.62466 | -81.3449
0.62498 | -81.3372
0.6253 | -81.3379
0.62563 | -81.3312
0.62595 | -81.3206
0.62627 | -81.3081
0.62659 | -81.2916
0.62691 | -81.2727
0.62723 | -81.2583
0.62755 | -81.236
0.62788 | -81.2126
0.6282 | -81.1981
0.62852 | -81.1872
0.62884 | -81.1776
0.62916 | -81.1623
0.62948 | -81.1485
0.6298 | -81.1416
0.63013 | -81.1376
0.63045 | -81.1214
0.63077 | -81.1078
0.63109 | -81.0884
0.63141 | -81.074
0.63173 | -81.0639
0.63205 | -81.0545
0.63237 | -81.048
0.6327 | -81.0472
0.63302 | -81.0434
0.63334 | -81.037
0.63366 | -81.0351
0.63398 | -81.0233
0.6343 | -81.0257
0.63462 | -81.0226
0.63494 | -81.0248
0.63527 | -81.0315
0.63559 | -81.0318
0.63591 | -81.0389
0.63623 | -81.0501
0.63655 | -81.0555
0.63687 | -81.0475
0.63719 | -81.0534
0.63752 | -81.0604
0.63784 | -81.0636
0.63816 | -81.0657
0.63848 | -81.0736
0.6388 | -81.0813
0.63912 | -81.0922
0.63944 | -81.1149
0.63977 | -81.1228
0.64009 | -81.136
0.64041 | -81.1503
0.64073 | -81.1677
0.64105 | -81.1791
0.64137 | -81.1917
0.64169 | -81.2062
0.64201 | -81.2201
0.64234 | -81.2395
0.64266 | -81.255
0.64298 | -81.269
0.6433 | -81.2913
0.64362 | -81.3143
0.64394 | -81.3344
0.64426 | -81.3557
0.64458 | -81.3766
0.64491 | -81.3953
0.64523 | -81.4169
0.64555 | -81.4439
0.64587 | -81.4649
0.64619 | -81.4756
0.64651 | -81.4921
0.64683 | -81.5145
0.64716 | -81.5308
0.64748 | -81.5473
0.6478 | -81.5671
0.64812 | -81.5887
0.64844 | -81.6125
0.64876 | -81.636
0.64908 | -81.6612
0.64941 | -81.6777
0.64973 | -81.6973
0.65005 | -81.7207
0.65037 | -81.7413
0.65069 | -81.7623
0.65101 | -81.7812
0.65133 | -81.7982
0.65165 | -81.8181
0.65198 | -81.8383
0.6523 | -81.8607
0.65262 | -81.8933
0.65294 | -81.9152
0.65326 | -81.9327
0.65358 | -81.9468
0.6539 | -81.9671
0.65422 | -81.9888
0.65455 | -82.0079
0.65487 | -82.0247
0.65519 | -82.0432
0.65551 | -82.0705
0.65583 | -82.0907
0.65615 | -82.1041
0.65647 | -82.1194
0.6568 | -82.1381
0.65712 | -82.1478
0.65744 | -82.1616
0.65776 | -82.1613
0.65808 | -82.1469
0.6584 | -82.1531
0.65872 | -82.1631
0.65904 | -82.1853
0.65937 | -82.2038
0.65969 | -82.2083
0.66001 | -82.2246
0.66033 | -82.241
0.66065 | -82.2436
0.66097 | -82.2473
0.66129 | -82.2477
0.66161 | -82.2459
0.66194 | -82.2449
0.66226 | -82.2456
0.66258 | -82.2439
0.6629 | -82.2477
0.66322 | -82.2494
0.66354 | -82.2438
0.66387 | -82.2415
0.66419 | -82.2507
0.66451 | -82.2621
0.66483 | -82.2575
0.66515 | -82.2578
0.66547 | -82.2539
0.66579 | -82.2505
0.66611 | -82.2438
0.66643 | -82.2318
0.66676 | -82.2234
0.66708 | -82.2207
0.6674 | -82.2005
0.66772 | -82.1833
0.66804 | -82.1687
0.66836 | -82.1581
0.66868 | -82.144
0.66901 | -82.1351
0.66933 | -82.1235
0.66965 | -82.1071
0.66997 | -82.0906
0.67029 | -82.0764
0.67061 | -82.0665
0.67093 | -82.0453
0.67126 | -82.0204
0.67158 | -82.0006
0.6719 | -81.981
0.67222 | -81.967
0.67254 | -81.9471
0.67286 | -81.9159
0.67318 | -81.9085
0.67351 | -81.8919
0.67383 | -81.8706
0.67415 | -81.8512
0.67447 | -81.8306
0.67479 | -81.8187
0.67511 | -81.8014
0.67543 | -81.7778
0.67575 | -81.755
0.67607 | -81.7342
0.6764 | -81.699
0.67672 | -81.6745
0.67704 | -81.6591
0.67736 | -81.6471
0.67768 | -81.6299
0.678 | -81.6095
0.67832 | -81.5882
0.67865 | -81.5684
0.67897 | -81.5503
0.67929 | -81.5321
0.67961 | -81.5095
0.67993 | -81.4842
0.68025 | -81.4717
0.68057 | -81.4511
0.6809 | -81.4318
0.68122 | -81.4133
0.68154 | -81.3899
0.68186 | -81.3683
0.68218 | -81.3518
0.6825 | -81.335
0.68282 | -81.3127
0.68314 | -81.2991
0.68347 | -81.2849
0.68379 | -81.2689
0.68411 | -81.2539
0.68443 | -81.2364
0.68475 | -81.2142
0.68507 | -81.1964
0.68539 | -81.1821
0.68571 | -81.1609
0.68604 | -81.151
0.68636 | -81.1368
0.68668 | -81.1239
0.687 | -81.1078
0.68732 | -81.0907
0.68764 | -81.0747
0.68796 | -81.0663
0.68829 | -81.0553
0.68861 | -81.045
0.68893 | -81.0358
0.68925 | -81.0255
0.68957 | -81.0183
0.68989 | -81.0122
0.69021 | -81.0044
0.69054 | -80.996
0.69086 | -80.9877
0.69118 | -80.9808
0.6915 | -80.9737
0.69182 | -80.9659
0.69214 | -80.956
0.69246 | -80.9472
0.69278 | -80.9328
0.69311 | -80.9233
0.69343 | -80.9219
0.69375 | -80.9183
0.69407 | -80.9191
0.69439 | -80.921
0.69471 | -80.9187
0.69503 | -80.9306
0.69535 | -80.9307
0.69568 | -80.9286
0.696 | -80.9299
0.69632 | -80.933
0.69664 | -80.9373
0.69696 | -80.9447
0.69728 | -80.9537
0.6976 | -80.956
0.69793 | -80.9604
0.69825 | -80.9651
0.69857 | -80.9731
0.69889 | -80.987
0.69921 | -80.9971
0.69953 | -81.0074
0.69985 | -81.0186
0.70017 | -81.0324
0.7005 | -81.0474
0.70082 | -81.0544
0.70114 | -81.0659
0.70146 | -81.0791
0.70178 | -81.0911
0.7021 | -81.1053
0.70242 | -81.1221
0.70274 | -81.1389
0.70307 | -81.1568
0.70339 | -81.1749
0.70371 | -81.1986
0.70403 | -81.2227
0.70435 | -81.2502
0.70467 | -81.2697
0.705 | -81.2888
0.70532 | -81.3094
0.70564 | -81.3378
0.70596 | -81.3687
0.70628 | -81.3891
0.7066 | -81.4084
0.70692 | -81.4275
0.70724 | -81.4463
0.70756 | -81.4664
0.70789 | -81.485
0.70821 | -81.4994
0.70853 | -81.52
0.70885 | -81.5423
0.70917 | -81.564
0.70949 | -81.588
0.70981 | -81.614
0.71014 | -81.6395
0.71046 | -81.6645
0.71078 | -81.6901
0.7111 | -81.7138
0.71142 | -81.7393
0.71174 | -81.7669
0.71206 | -81.7951
0.71239 | -81.8196
0.71271 | -81.8442
0.71303 | -81.8672
0.71335 | -81.8905
0.71367 | -81.9181
0.71399 | -81.9437
0.71431 | -81.9636
0.71464 | -81.9749
0.71496 | -81.9842
0.71528 | -82.0008
0.7156 | -82.0173
0.71592 | -82.0352
0.71624 | -82.0617
0.71656 | -82.0886
0.71688 | -82.1118
0.7172 | -82.1313
0.71753 | -82.1539
0.71785 | -82.1716
0.71817 | -82.1946
0.71849 | -82.2171
0.71881 | -82.2347
0.71913 | -82.2555
0.71945 | -82.2676
0.71978 | -82.2894
0.7201 | -82.3167
0.72042 | -82.3352
0.72074 | -82.355
0.72106 | -82.3683
0.72138 | -82.3905
0.7217 | -82.4134
0.72203 | -82.421
0.72235 | -82.4393
0.72267 | -82.4514
0.72299 | -82.4596
0.72331 | -82.471
0.72363 | -82.4816
0.72395 | -82.4887
0.72427 | -82.497
0.7246 | -82.5104
0.72492 | -82.5225
0.72524 | -82.5334
0.72556 | -82.5614
0.72588 | -82.57
0.7262 | -82.5735
0.72652 | -82.5799
0.72684 | -82.5786
0.72717 | -82.575
0.72749 | -82.5755
0.72781 | -82.5794
0.72813 | -82.5822
0.72845 | -82.5827
0.72877 | -82.5849
0.72909 | -82.5931
0.72942 | -82.6025
0.72974 | -82.6066
0.73006 | -82.6078
0.73038 | -82.6081
0.7307 | -82.6107
0.73102 | -82.6179
0.73134 | -82.6226
0.73167 | -82.622
0.73199 | -82.6216
0.73231 | -82.6257
0.73263 | -82.6322
0.73295 | -82.6405
0.73327 | -82.6332
0.73359 | -82.6245
0.73391 | -82.6218
0.73424 | -82.6202
0.73456 | -82.6221
0.73488 | -82.6093
0.7352 | -82.6068
0.73552 | -82.6076
0.73584 | -82.6048
0.73616 | -82.5963
0.73648 | -82.5879
0.73681 | -82.5844
0.73713 | -82.5833
0.73745 | -82.5875
0.73777 | -82.5866
0.73809 | -82.584
0.73841 | -82.5819
0.73873 | -82.5783
0.73906 | -82.5742
0.73938 | -82.5716
0.7397 | -82.5654
0.74002 | -82.5614
0.74034 | -82.5575
0.74066 | -82.5567
0.74098 | -82.5594
0.7413 | -82.5599
0.74163 | -82.5673
0.74195 | -82.57
0.74227 | -82.5727
0.74259 | -82.5805
0.74291 | -82.584
0.74323 | -82.5982
0.74355 | -82.6173
0.74387 | -82.6183
0.7442 | -82.6175
0.74452 | -82.6171
0.74484 | -82.61
0.74516 | -82.609
0.74548 | -82.61
0.7458 | -82.6107
0.74613 | -82.612
0.74645 | -82.6177
0.74677 | -82.6182
0.74709 | -82.6173
0.74741 | -82.6197
0.74773 | -82.6202
0.74805 | -82.6163
0.74837 | -82.6228
0.74869 | -82.6266
0.74902 | -82.6173
0.74934 | -82.6122
0.74966 | -82.6143
0.74998 | -82.6179
0.7503 | -82.6204
0.75062 | -82.6239
0.75094 | -82.625
0.75127 | -82.6271
0.75159 | -82.6311
0.75191 | -82.6431
0.75223 | -82.6535
0.75255 | -82.6564
0.75287 | -82.6622
0.75319 | -82.6671
0.75352 | -82.6712
0.75384 | -82.6756
0.75416 | -82.6799
0.75448 | -82.6929
0.7548 | -82.6954
0.75512 | -82.6921
0.75544 | -82.6927
0.75577 | -82.6959
0.75609 | -82.6932
0.75641 | -82.6959
0.75673 | -82.6997
0.75705 | -82.7058
0.75737 | -82.7081
0.75769 | -82.7016
0.75801 | -82.6974
0.75833 | -82.6988
0.75866 | -82.7006
0.75898 | -82.7027
0.7593 | -82.7074
0.75962 | -82.7197
0.75994 | -82.7204
0.76026 | -82.7174
0.76058 | -82.711
0.76091 | -82.7076
0.76123 | -82.7087
0.76155 | -82.7034
0.76187 | -82.6994
0.76219 | -82.7005
0.76251 | -82.6989
0.76283 | -82.6996
0.76316 | -82.7
0.76348 | -82.7025
0.7638 | -82.7062
0.76412 | -82.7105
0.76444 | -82.7127
0.76476 | -82.7142
0.76508 | -82.7149
0.7654 | -82.7166
0.76573 | -82.717
0.76605 | -82.7174
0.76637 | -82.7232
0.76669 | -82.7411
0.76701 | -82.7613
0.76733 | -82.7699
0.76765 | -82.7738
0.76797 | -82.7779
0.7683 | -82.7811
0.76862 | -82.7823
0.76894 | -82.7823
0.76926 | -82.7887
0.76958 | -82.7903
0.7699 | -82.7899
0.77022 | -82.7931
0.77055 | -82.7953
0.77087 | -82.8014
0.77119 | -82.81
0.77151 | -82.8125
0.77183 | -82.8209
0.77215 | -82.8284
0.77247 | -82.8351
0.7728 | -82.8403
0.77312 | -82.8439
0.77344 | -82.8489
0.77376 | -82.8512
0.77408 | -82.8593
0.7744 | -82.8665
0.77472 | -82.8735
0.77504 | -82.8815
0.77537 | -82.889
0.77569 | -82.8982
0.77601 | -82.9052
0.77633 | -82.9139
0.77665 | -82.9249
0.77697 | -82.9375
0.77729 | -82.9463
0.77761 | -82.9536
0.77794 | -82.9667
0.77826 | -82.9802
0.77858 | -83.004
0.7789 | -83.0234
0.77922 | -83.0367
0.77954 | -83.0454
0.77986 | -83.0572
0.78019 | -83.0712
0.78051 | -83.0775
0.78083 | -83.0703
0.78115 | -83.0712
0.78147 | -83.0811
0.78179 | -83.0948
0.78211 | -83.1085
0.78243 | -83.1217
0.78276 | -83.1335
0.78308 | -83.1538
0.7834 | -83.1697
0.78372 | -83.1814
0.78404 | -83.1958
0.78436 | -83.2058
0.78468 | -83.2128
0.78501 | -83.2248
0.78533 | -83.2383
0.78565 | -83.2511
0.78597 | -83.2519
0.78629 | -83.2597
0.78661 | -83.2821
0.78693 | -83.2982
0.78726 | -83.299
0.78758 | -83.3106
0.7879 | -83.3153
0.78822 | -83.3302
0.78854 | -83.3488
0.78886 | -83.3472
0.78918 | -83.3492
0.7895 | -83.3537
0.78982 | -83.3617
0.79015 | -83.3684
0.79047 | -83.3711
0.79079 | -83.3712
0.79111 | -83.3629
0.79143 | -83.358
0.79175 | -83.3483
0.79207 | -83.3395
0.7924 | -83.3364
0.79272 | -83.333
0.79304 | -83.3286
0.79336 | -83.3289
0.79368 | -83.3262
0.794 | -83.3222
0.79432 | -83.3153
0.79465 | -83.3013
0.79497 | -83.2864
0.79529 | -83.2732
0.79561 | -83.2601
0.79593 | -83.2484
0.79625 | -83.2355
0.79657 | -83.2193
0.7969 | -83.198
0.79722 | -83.1625
0.79754 | -83.1415
0.79786 | -83.1177
0.79818 | -83.0923
0.7985 | -83.0702
0.79882 | -83.0501
0.79914 | -83.0373
0.79946 | -83.0184
0.79979 | -82.996
0.80011 | -82.9736
0.80043 | -82.9518
0.80075 | -82.9292
0.80107 | -82.9087
0.80139 | -82.8888
0.80171 | -82.8647
0.80204 | -82.8447
0.80236 | -82.8281
0.80268 | -82.8007
0.803 | -82.7777
0.80332 | -82.7678
0.80364 | -82.751
0.80396 | -82.7238
0.80429 | -82.6961
0.80461 | -82.6628
0.80493 | -82.6309
0.80525 | -82.6073
0.80557 | -82.5901
0.80589 | -82.5686
0.80621 | -82.5462
0.80654 | -82.5327
0.80686 | -82.5119
0.80718 | -82.4896
0.8075 | -82.4693
0.80782 | -82.4522
0.80814 | -82.4318
0.80846 | -82.4109
0.80878 | -82.3925
0.8091 | -82.3743
0.80943 | -82.352
0.80975 | -82.3409
0.81007 | -82.3286
0.81039 | -82.314
0.81071 | -82.3008
0.81103 | -82.2929
0.81135 | -82.285
0.81168 | -82.2796
0.812 | -82.2695
0.81232 | -82.2596
0.81264 | -82.2576
0.81296 | -82.2545
0.81328 | -82.2477
0.8136 | -82.243
0.81393 | -82.2377
0.81425 | -82.2321
0.81457 | -82.2194
0.81489 | -82.2137
0.81521 | -82.2145
0.81553 | -82.217
0.81585 | -82.2294
0.81617 | -82.2349
0.8165 | -82.2378
0.81682 | -82.2388
0.81714 | -82.2458
0.81746 | -82.2507
0.81778 | -82.2525
0.8181 | -82.2591
0.81842 | -82.2653
0.81874 | -82.2692
0.81907 | -82.2708
0.81939 | -82.2772
0.81971 | -82.2861
0.82003 | -82.2939
0.82035 | -82.304
0.82067 | -82.3162
0.82099 | -82.3253
0.82132 | -82.334
0.82164 | -82.3463
0.82196 | -82.358
0.82228 | -82.3644
0.8226 | -82.3718
0.82292 | -82.3832
0.82324 | -82.3991
0.82356 | -82.4148
0.82389 | -82.4305
0.82421 | -82.4442
0.82453 | -82.4594
0.82485 | -82.4748
0.82517 | -82.4913
0.82549 | -82.5068
0.82581 | -82.5183
0.82614 | -82.528
0.82646 | -82.5374
0.82678 | -82.5462
0.8271 | -82.5619
0.82742 | -82.5736
0.82774 | -82.584
0.82806 | -82.5992
0.82839 | -82.6155
0.82871 | -82.6316
0.82903 | -82.6499
0.82935 | -82.6666
0.82967 | -82.6868
0.82999 | -82.7062
0.83031 | -82.7249
0.83063 | -82.7422
0.83096 | -82.7568
0.83128 | -82.77
0.8316 | -82.7819
0.83192 | -82.7902
0.83224 | -82.8017
0.83256 | -82.8173
0.83288 | -82.8316
0.8332 | -82.8441
0.83353 | -82.858
0.83385 | -82.8748
0.83417 | -82.8809
0.83449 | -82.8901
0.83481 | -82.9025
0.83513 | -82.9148
0.83545 | -82.9251
0.83578 | -82.9356
0.8361 | -82.9496
0.83642 | -82.963
0.83674 | -82.9769
0.83706 | -82.9912
0.83738 | -83.0084
0.8377 | -83.0339
0.83803 | -83.0515
0.83835 | -83.0609
0.83867 | -83.0661
0.83899 | -83.0708
0.83931 | -83.0773
0.83963 | -83.0837
0.83995 | -83.0868
0.84027 | -83.0931
0.84059 | -83.0989
0.84092 | -83.1035
0.84124 | -83.1123
0.84156 | -83.1235
0.84188 | -83.1272
0.8422 | -83.132
0.84252 | -83.1373
0.84284 | -83.1546
0.84317 | -83.1652
0.84349 | -83.1734
0.84381 | -83.1773
0.84413 | -83.1847
0.84445 | -83.199
0.84477 | -83.2044
0.84509 | -83.2086
0.84542 | -83.2137
0.84574 | -83.2196
0.84606 | -83.2253
0.84638 | -83.2372
0.8467 | -83.2483
0.84702 | -83.2548
0.84734 | -83.2597
0.84767 | -83.2641
0.84799 | -83.2687
0.84831 | -83.2753
0.84863 | -83.2822
0.84895 | -83.2918
0.84927 | -83.3056
0.84959 | -83.3164
0.84991 | -83.323
0.85023 | -83.3331
0.85056 | -83.3422
0.85088 | -83.3495
0.8512 | -83.3578
0.85152 | -83.3675
0.85184 | -83.375
0.85216 | -83.3826
0.85249 | -83.3922
0.85281 | -83.4095
0.85313 | -83.4232
0.85345 | -83.4343
0.85377 | -83.4436
0.85409 | -83.461
0.85441 | -83.4793
0.85473 | -83.5014
0.85506 | -83.5243
0.85538 | -83.5391
0.8557 | -83.556
0.85602 | -83.5728
0.85634 | -83.5875
0.85666 | -83.6038
0.85698 | -83.6144
0.8573 | -83.6256
0.85763 | -83.6321
0.85795 | -83.6348
0.85827 | -83.6456
0.85859 | -83.6596
0.85891 | -83.6732
0.85923 | -83.6874
0.85955 | -83.7014
0.85987 | -83.7173
0.8602 | -83.7346
0.86052 | -83.7514
0.86084 | -83.7692
0.86116 | -83.7854
0.86148 | -83.8015
0.8618 | -83.8183
0.86212 | -83.8342
0.86245 | -83.8531
0.86277 | -83.8709
0.86309 | -83.8965
0.86341 | -83.9185
0.86373 | -83.9303
0.86405 | -83.9443
0.86437 | -83.9561
0.86469 | -83.9681
0.86502 | -83.9834
0.86534 | -83.9963
0.86566 | -84.0108
0.86598 | -84.0281
0.8663 | -84.0462
0.86662 | -84.058
0.86694 | -84.0679
0.86727 | -84.0801
0.86759 | -84.0864
0.86791 | -84.0936
0.86823 | -84.1063
0.86855 | -84.11
0.86887 | -84.1201
0.86919 | -84.1301
0.86952 | -84.1353
0.86984 | -84.1417
0.87016 | -84.1434
0.87048 | -84.1411
0.8708 | -84.1457
0.87112 | -84.1527
0.87144 | -84.1488
0.87176 | -84.1525
0.87209 | -84.1496
0.87241 | -84.1481
0.87273 | -84.1422
0.87305 | -84.1292
0.87337 | -84.1289
0.87369 | -84.1361
0.87401 | -84.1373
0.87433 | -84.1285
0.87466 | -84.1257
0.87498 | -84.1232
0.8753 | -84.1197
0.87562 | -84.1086
0.87594 | -84.1
0.87626 | -84.0895
0.87658 | -84.0774
0.87691 | -84.0723
0.87723 | -84.0628
0.87755 | -84.0533
0.87787 | -84.0437
0.87819 | -84.0375
0.87851 | -84.0347
0.87883 | -84.02
0.87916 | -84.0063
0.87948 | -83.9983
0.8798 | -83.9892
0.88012 | -83.9714
0.88044 | -83.9495
0.88076 | -83.9412
0.88108 | -83.935
0.8814 | -83.9242
0.88172 | -83.9154
0.88205 | -83.9115
0.88237 | -83.9063
0.88269 | -83.8938
0.88301 | -83.877
0.88333 | -83.8647
0.88365 | -83.8521
0.88397 | -83.8432
0.8843 | -83.8294
0.88462 | -83.8139
0.88494 | -83.8086
0.88526 | -83.7968
0.88558 | -83.7876
0.8859 | -83.7828
0.88622 | -83.7822
0.88655 | -83.7785
0.88687 | -83.7751
0.88719 | -83.7648
0.88751 | -83.7596
0.88783 | -83.7568
0.88815 | -83.7451
0.88847 | -83.73
0.8888 | -83.7178
0.88912 | -83.7055
0.88944 | -83.6975
0.88976 | -83.7001
0.89008 | -83.6928
0.8904 | -83.6883
0.89072 | -83.6826
0.89104 | -83.6739
0.89136 | -83.669
0.89169 | -83.6648
0.89201 | -83.6623
0.89233 | -83.665
0.89265 | -83.6655
0.89297 | -83.6624
0.89329 | -83.6596
0.89362 | -83.6574
0.89394 | -83.6563
0.89426 | -83.6549
0.89458 | -83.6567
0.8949 | -83.6595
0.89522 | -83.6619
0.89554 | -83.6586
0.89586 | -83.6576
0.89619 | -83.6618
0.89651 | -83.6666
0.89683 | -83.6696
0.89715 | -83.6773
0.89747 | -83.6866
0.89779 | -83.6935
0.89811 | -83.7004
0.89843 | -83.7075
0.89876 | -83.7144
0.89908 | -83.7334
0.8994 | -83.7399
0.89972 | -83.7491
0.90004 | -83.7574
0.90036 | -83.7628
0.90068 | -83.771
0.901 | -83.7838
0.90133 | -83.7962
0.90165 | -83.8092
0.90197 | -83.8226
0.90229 | -83.836
0.90261 | -83.8504
0.90293 | -83.8669
0.90325 | -83.8822
0.90358 | -83.8954
0.9039 | -83.9081
0.90422 | -83.9241
0.90454 | -83.9396
0.90486 | -83.958
0.90518 | -83.9778
0.9055 | -83.9927
0.90582 | -84.008
0.90615 | -84.0284
0.90647 | -84.047
0.90679 | -84.0669
0.90711 | -84.0866
0.90743 | -84.1027
0.90775 | -84.12
0.90807 | -84.1364
0.9084 | -84.1586
0.90872 | -84.1677
0.90904 | -84.1778
0.90936 | -84.1977
0.90968 | -84.2207
0.91 | -84.2437
0.91032 | -84.2712
0.91065 | -84.2907
0.91097 | -84.2996
0.91129 | -84.3136
0.91161 | -84.3324
0.91193 | -84.3508
0.91225 | -84.3712
0.91257 | -84.3922
0.91289 | -84.412
0.91322 | -84.4304
0.91354 | -84.4531
0.91386 | -84.4793
0.91418 | -84.5046
0.9145 | -84.5256
0.91482 | -84.5437
0.91514 | -84.5606
0.91546 | -84.579
0.91579 | -84.5992
0.91611 | -84.6154
0.91643 | -84.6302
0.91675 | -84.6458
0.91707 | -84.6575
0.91739 | -84.6688
0.91771 | -84.6733
0.91804 | -84.6881
0.91836 | -84.7068
0.91868 | -84.7285
0.919 | -84.7407
0.91932 | -84.7555
0.91964 | -84.7691
0.91996 | -84.7816
0.92029 | -84.8019
0.92061 | -84.8109
0.92093 | -84.8152
0.92125 | -84.8207
0.92157 | -84.8252
0.92189 | -84.8297
0.92221 | -84.8346
0.92253 | -84.837
0.92285 | -84.8418
0.92318 | -84.8444
0.9235 | -84.838
0.92382 | -84.8279
0.92414 | -84.8252
0.92446 | -84.8236
0.92478 | -84.8207
0.9251 | -84.816
0.92543 | -84.8137
0.92575 | -84.809
0.92607 | -84.804
0.92639 | -84.8016
0.92671 | -84.795
0.92703 | -84.7783
0.92735 | -84.772
0.92768 | -84.7657
0.928 | -84.7557
0.92832 | -84.743
0.92864 | -84.7308
0.92896 | -84.72
0.92928 | -84.7101
0.9296 | -84.6992
0.92993 | -84.6932
0.93025 | -84.6849
0.93057 | -84.6751
0.93089 | -84.6639
0.93121 | -84.65
0.93153 | -84.6358
0.93185 | -84.6212
0.93217 | -84.6044
0.93249 | -84.5894
0.93282 | -84.5741
0.93314 | -84.5609
0.93346 | -84.5531
0.93378 | -84.5391
0.9341 | -84.524
0.93442 | -84.5082
0.93475 | -84.494
0.93507 | -84.4716
0.93539 | -84.4543
0.93571 | -84.4404
0.93603 | -84.4309
0.93635 | -84.4198
0.93667 | -84.4139
0.93699 | -84.3996
0.93732 | -84.3886
0.93764 | -84.3738
0.93796 | -84.3605
0.93828 | -84.3438
0.9386 | -84.327
0.93892 | -84.3084
0.93924 | -84.2987
0.93956 | -84.2849
0.93989 | -84.2719
0.94021 | -84.2616
0.94053 | -84.2443
0.94085 | -84.2301
0.94117 | -84.2164
0.94149 | -84.2032
0.94181 | -84.1908
0.94213 | -84.1795
0.94246 | -84.1722
0.94278 | -84.1632
0.9431 | -84.1495
0.94342 | -84.1365
0.94374 | -84.1319
0.94406 | -84.1265
0.94438 | -84.1167
0.94471 | -84.1064
0.94503 | -84.097
0.94535 | -84.0928
0.94567 | -84.079
0.94599 | -84.0653
0.94631 | -84.0546
0.94663 | -84.0357
0.94695 | -84.0233
0.94728 | -84.018
0.9476 | -84.01
0.94792 | -84.0017
0.94824 | -83.9951
0.94856 | -83.9908
0.94888 | -83.9903
0.9492 | -83.9783
0.94953 | -83.964
0.94985 | -83.9564
0.95017 | -83.9504
0.95049 | -83.9404
0.95081 | -83.9301
0.95113 | -83.9238
0.95145 | -83.9121
0.95178 | -83.9016
0.9521 | -83.8937
0.95242 | -83.8906
0.95274 | -83.8894
0.95306 | -83.8856
0.95338 | -83.8831
0.9537 | -83.8783
0.95402 | -83.8709
0.95435 | -83.8601
0.95467 | -83.8584
0.95499 | -83.854
0.95531 | -83.8476
0.95563 | -83.8405
0.95595 | -83.8359
0.95627 | -83.8289
0.95659 | -83.8266
0.95692 | -83.8235
0.95724 | -83.8194
0.95756 | -83.8138
0.95788 | -83.807
0.9582 | -83.7958
0.95852 | -83.7874
0.95884 | -83.784
0.95917 | -83.7802
0.95949 | -83.7762
0.95981 | -83.7703
0.96013 | -83.7655
0.96045 | -83.7648
0.96077 | -83.7611
0.96109 | -83.7555
0.96142 | -83.7449
0.96174 | -83.7346
0.96206 | -83.7292
0.96238 | -83.7246
0.9627 | -83.7171
0.96302 | -83.712
0.96334 | -83.7043
0.96366 | -83.6937
0.96399 | -83.6907
0.96431 | -83.6864
0.96463 | -83.6792
0.96495 | -83.6767
0.96527 | -83.672
0.96559 | -83.6679
0.96591 | -83.6658
0.96623 | -83.6544
0.96656 | -83.6487
0.96688 | -83.6418
0.9672 | -83.6349
0.96752 | -83.6302
0.96784 | -83.6251
0.96816 | -83.6216
0.96848 | -83.6167
0.96881 | -83.6112
0.96913 | -83.6092
0.96945 | -83.6075
0.96977 | -83.6051
0.97009 | -83.6013
0.97041 | -83.5968
0.97073 | -83.593
0.97106 | -83.5946
0.97138 | -83.5962
0.9717 | -83.5943
0.97202 | -83.5921
0.97234 | -83.5879
0.97266 | -83.585
0.97298 | -83.5861
0.9733 | -83.5904
0.97362 | -83.5907
0.97395 | -83.5918
0.97427 | -83.5901
0.97459 | -83.5884
0.97491 | -83.5948
0.97523 | -83.6027
0.97555 | -83.6043
0.97588 | -83.6018
0.9762 | -83.5986
0.97652 | -83.5959
0.97684 | -83.5962
0.97716 | -83.6029
0.97748 | -83.6061
0.9778 | -83.6067
0.97812 | -83.6074
0.97845 | -83.6043
0.97877 | -83.5963
0.97909 | -83.5873
0.97941 | -83.5861
0.97973 | -83.5875
0.98005 | -83.5906
0.98037 | -83.5922
0.98069 | -83.5925
0.98102 | -83.5943
0.98134 | -83.5974
0.98166 | -83.6003
0.98198 | -83.6025
0.9823 | -83.6043
0.98262 | -83.6089
0.98294 | -83.614
0.98326 | -83.617
0.98359 | -83.6232
0.98391 | -83.6307
0.98423 | -83.6326
0.98455 | -83.6353
0.98487 | -83.6398
0.98519 | -83.645
0.98552 | -83.6506
0.98584 | -83.6559
0.98616 | -83.6611
0.98648 | -83.6649
0.9868 | -83.6783
0.98712 | -83.6837
0.98744 | -83.691
0.98776 | -83.6958
0.98808 | -83.6997
0.98841 | -83.7015
0.98873 | -83.7069
0.98905 | -83.7144
0.98937 | -83.7213
0.98969 | -83.7274
0.99001 | -83.7358
0.99033 | -83.7463
0.99066 | -83.7598
0.99098 | -83.7684
0.9913 | -83.7754
0.99162 | -83.7817
0.99194 | -83.7912
0.99226 | -83.7983
0.99258 | -83.8058
0.99291 | -83.8141
0.99323 | -83.8197
0.99355 | -83.8263
0.99387 | -83.8332
0.99419 | -83.8437
0.99451 | -83.8509
0.99483 | -83.8571
0.99515 | -83.8655
0.99548 | -83.8704
0.9958 | -83.8798
0.99612 | -83.8898
0.99644 | -83.8972
0.99676 | -83.8944
0.99708 | -83.8867
0.9974 | -83.89
0.99772 | -83.8963
0.99805 | -83.9043
0.99837 | -83.9106
0.99869 | -83.9185
0.99901 | -83.9279
0.99933 | -83.934
0.99965 | -83.9389
0.99997 | -83.9416
1.0003 | -83.9462
1.0006 | -83.951
1.0009 | -83.9501
1.0013 | -83.9515
1.0016 | -83.9545
1.0019 | -83.9556
1.0022 | -83.9565
1.0025 | -83.9596
1.0029 | -83.9652
1.0032 | -83.9706
1.0035 | -83.9678
1.0038 | -83.9699
1.0042 | -83.9707
1.0045 | -83.9727
1.0048 | -83.9768
1.0051 | -83.9795
1.0054 | -83.9806
1.0058 | -83.9787
1.0061 | -83.978
1.0064 | -83.9794
1.0067 | -83.9798
1.007 | -83.9818
1.0074 | -83.9794
1.0077 | -83.9779
1.008 | -83.977
1.0083 | -83.9768
1.0087 | -83.9769
1.009 | -83.9762
1.0093 | -83.9742
1.0096 | -83.9738
1.0099 | -83.971
1.0103 | -83.9652
1.0106 | -83.9587
1.0109 | -83.956
1.0112 | -83.9545
1.0115 | -83.9547
1.0119 | -83.9546
1.0122 | -83.9522
1.0125 | -83.9488
1.0128 | -83.9471
1.0131 | -83.9454
1.0135 | -83.9423
1.0138 | -83.9413
1.0141 | -83.9418
1.0144 | -83.9439
1.0148 | -83.9474
1.0151 | -83.9463
1.0154 | -83.9505
1.0157 | -83.9484
1.016 | -83.943
1.0164 | -83.9398
1.0167 | -83.9378
1.017 | -83.9347
1.0173 | -83.9306
1.0176 | -83.9275
1.018 | -83.9291
1.0183 | -83.9246
1.0186 | -83.9174
1.0189 | -83.9136
1.0193 | -83.9136
1.0196 | -83.911
1.0199 | -83.9131
1.0202 | -83.9124
1.0205 | -83.9149
1.0209 | -83.9172
1.0212 | -83.9156
1.0215 | -83.9087
1.0218 | -83.9071
1.0221 | -83.9044
1.0225 | -83.9032
1.0228 | -83.8998
1.0231 | -83.8986
1.0234 | -83.9025
1.0238 | -83.9008
1.0241 | -83.899
1.0244 | -83.8975
1.0247 | -83.896
1.025 | -83.8973
1.0254 | -83.8996
1.0257 | -83.9015
1.026 | -83.9006
1.0263 | -83.9017
1.0266 | -83.8992
1.027 | -83.9
1.0273 | -83.903
1.0276 | -83.9113
1.0279 | -83.9162
1.0283 | -83.9163
1.0286 | -83.915
1.0289 | -83.9167
1.0292 | -83.9181
1.0295 | -83.9212
1.0299 | -83.9243
1.0302 | -83.9273
1.0305 | -83.9306
1.0308 | -83.9335
1.0311 | -83.9423
1.0315 | -83.9511
1.0318 | -83.9528
1.0321 | -83.9489
1.0324 | -83.9553
1.0328 | -83.9595
1.0331 | -83.9654
1.0334 | -83.9746
1.0337 | -83.979
1.034 | -83.9822
1.0344 | -83.986
1.0347 | -83.9909
1.035 | -83.989
1.0353 | -83.9862
1.0356 | -83.9842
1.036 | -83.9878
1.0363 | -84.0014
1.0366 | -84.0066
1.0369 | -84.0133
1.0372 | -84.0189
1.0376 | -84.022
1.0379 | -84.0238
1.0382 | -84.0292
1.0385 | -84.0328
1.0389 | -84.0394
1.0392 | -84.0465
1.0395 | -84.0523
1.0398 | -84.0595
1.0401 | -84.0646
1.0405 | -84.0663
1.0408 | -84.0687
1.0411 | -84.0735
1.0414 | -84.0799
1.0417 | -84.087
1.0421 | -84.0938
1.0424 | -84.1012
1.0427 | -84.1086
1.043 | -84.115
1.0434 | -84.1232
1.0437 | -84.1282
1.044 | -84.1314
1.0443 | -84.1353
1.0446 | -84.1422
1.045 | -84.1494
1.0453 | -84.1552
1.0456 | -84.1643
1.0459 | -84.1713
1.0462 | -84.1773
1.0466 | -84.1843
1.0469 | -84.1929
1.0472 | -84.2041
1.0475 | -84.2133
1.0479 | -84.2198
1.0482 | -84.2248
1.0485 | -84.2315
1.0488 | -84.2354
1.0491 | -84.2398
1.0495 | -84.2495
1.0498 | -84.2652
1.0501 | -84.2807
1.0504 | -84.2866
1.0507 | -84.2905
1.0511 | -84.2944
1.0514 | -84.2976
1.0517 | -84.3025
1.052 | -84.3083
1.0524 | -84.3123
1.0527 | -84.3187
1.053 | -84.3235
1.0533 | -84.3303
1.0536 | -84.3395
1.054 | -84.3485
1.0543 | -84.3546
1.0546 | -84.362
1.0549 | -84.3637
1.0552 | -84.3637
1.0556 | -84.3703
1.0559 | -84.3795
1.0562 | -84.3849
1.0565 | -84.3874
1.0568 | -84.3897
1.0572 | -84.3924
1.0575 | -84.3873
1.0578 | -84.3828
1.0581 | -84.3857
1.0585 | -84.3869
1.0588 | -84.3929
1.0591 | -84.3939
1.0594 | -84.3973
1.0597 | -84.4037
1.0601 | -84.4118
1.0604 | -84.4138
1.0607 | -84.4135
1.061 | -84.4116
1.0613 | -84.4075
1.0617 | -84.405
1.062 | -84.4087
1.0623 | -84.4077
1.0626 | -84.4083
1.063 | -84.4091
1.0633 | -84.4036
1.0636 | -84.3895
1.0639 | -84.3837
1.0642 | -84.381
1.0646 | -84.3746
1.0649 | -84.3663
1.0652 | -84.3636
1.0655 | -84.3565
1.0658 | -84.3512
1.0662 | -84.3453
1.0665 | -84.3421
1.0668 | -84.3329
1.0671 | -84.3264
1.0675 | -84.3228
1.0678 | -84.3231
1.0681 | -84.3227
1.0684 | -84.3167
1.0687 | -84.3116
1.0691 | -84.3171
1.0694 | -84.3273
1.0697 | -84.3243
1.07 | -84.3171
1.0703 | -84.3122
1.0707 | -84.3086
1.071 | -84.3014
1.0713 | -84.2935
1.0716 | -84.287
1.072 | -84.2801
1.0723 | -84.2731
1.0726 | -84.2663
1.0729 | -84.261
1.0732 | -84.2565
1.0736 | -84.2462
1.0739 | -84.2386
1.0742 | -84.228
1.0745 | -84.2079
1.0748 | -84.2005
1.0752 | -84.1897
1.0755 | -84.182
1.0758 | -84.1772
1.0761 | -84.1699
1.0764 | -84.1589
1.0768 | -84.1467
1.0771 | -84.1365
1.0774 | -84.1276
1.0777 | -84.1189
1.0781 | -84.1085
1.0784 | -84.099
1.0787 | -84.0932
1.079 | -84.0879
1.0793 | -84.0798
1.0797 | -84.0722
1.08 | -84.0653
1.0803 | -84.0579
1.0806 | -84.0485
1.0809 | -84.038
1.0813 | -84.0282
1.0816 | -84.0203
1.0819 | -84.0126
1.0822 | -84.0041
1.0826 | -83.9943
1.0829 | -83.9871
1.0832 | -83.9815
1.0835 | -83.9729
1.0838 | -83.967
1.0842 | -83.9591
1.0845 | -83.9537
1.0848 | -83.9429
1.0851 | -83.9335
1.0854 | -83.9225
1.0858 | -83.9131
1.0861 | -83.9066
1.0864 | -83.897
1.0867 | -83.8865
1.0871 | -83.875
1.0874 | -83.8665
1.0877 | -83.8584
1.088 | -83.8514
1.0883 | -83.8437
1.0887 | -83.8375
1.089 | -83.834
1.0893 | -83.8317
1.0896 | -83.8274
1.0899 | -83.821
1.0903 | -83.8115
1.0906 | -83.7996
1.0909 | -83.7933
1.0912 | -83.7872
1.0916 | -83.7814
1.0919 | -83.7828
1.0922 | -83.7821
1.0925 | -83.7776
1.0928 | -83.7721
1.0932 | -83.7679
1.0935 | -83.7663
1.0938 | -83.7709
1.0941 | -83.7676
1.0944 | -83.7646
1.0948 | -83.7622
1.0951 | -83.7596
1.0954 | -83.7585
1.0957 | -83.7606
1.0961 | -83.7597
1.0964 | -83.7608
1.0967 | -83.7605
1.097 | -83.7626
1.0973 | -83.7656
1.0977 | -83.7672
1.098 | -83.7692
1.0983 | -83.7716
1.0986 | -83.775
1.0989 | -83.7803
1.0993 | -83.7834
1.0996 | -83.7857
1.0999 | -83.7891
1.1002 | -83.7925
1.1005 | -83.7966
1.1009 | -83.8051
1.1012 | -83.8126
1.1015 | -83.8196
1.1018 | -83.8259
1.1022 | -83.8316
1.1025 | -83.8399
1.1028 | -83.8502
1.1031 | -83.8684
1.1034 | -83.881
1.1038 | -83.8925
1.1041 | -83.9031
1.1044 | -83.9132
1.1047 | -83.9237
1.105 | -83.9348
1.1054 | -83.9475
1.1057 | -83.9602
1.106 | -83.9725
1.1063 | -83.986
1.1067 | -84.002
1.107 | -84.023
1.1073 | -84.037
1.1076 | -84.0487
1.1079 | -84.0648
1.1083 | -84.0784
1.1086 | -84.0926
1.1089 | -84.1102
1.1092 | -84.1252
1.1095 | -84.1405
1.1099 | -84.1635
1.1102 | -84.1885
1.1105 | -84.2114
1.1108 | -84.2297
1.1112 | -84.2456
1.1115 | -84.2638
1.1118 | -84.2833
1.1121 | -84.3042
1.1124 | -84.3272
1.1128 | -84.3492
1.1131 | -84.3734
1.1134 | -84.3969
1.1137 | -84.4173
1.114 | -84.4361
1.1144 | -84.457
1.1147 | -84.4787
1.115 | -84.4999
1.1153 | -84.5211
1.1157 | -84.5419
1.116 | -84.571
1.1163 | -84.5974
1.1166 | -84.6191
1.1169 | -84.6421
1.1173 | -84.6652
1.1176 | -84.6879
1.1179 | -84.7113
1.1182 | -84.7371
1.1185 | -84.7637
1.1189 | -84.7884
1.1192 | -84.8147
1.1195 | -84.8446
1.1198 | -84.8755
1.1202 | -84.9017
1.1205 | -84.9287
1.1208 | -84.957
1.1211 | -84.9828
1.1214 | -85.0084
1.1218 | -85.0234
1.1221 | -85.0485
1.1224 | -85.0739
1.1227 | -85.0993
1.123 | -85.1221
1.1234 | -85.1454
1.1237 | -85.1666
1.124 | -85.1932
1.1243 | -85.222
1.1246 | -85.2492
1.125 | -85.2779
1.1253 | -85.3035
1.1256 | -85.3272
1.1259 | -85.3507
1.1263 | -85.3746
1.1266 | -85.3971
1.1269 | -85.4186
1.1272 | -85.4394
1.1275 | -85.4577
1.1279 | -85.476
1.1282 | -85.4928
1.1285 | -85.5061
1.1288 | -85.52
1.1291 | -85.5425
1.1295 | -85.5615
1.1298 | -85.5752
1.1301 | -85.5918
1.1304 | -85.6093
1.1308 | -85.61
1.1311 | -85.6241
1.1314 | -85.6446
1.1317 | -85.6624
1.132 | -85.6768
1.1324 | -85.6923
1.1327 | -85.7075
1.133 | -85.7219
1.1333 | -85.7363
1.1336 | -85.7506
1.134 | -85.7635
1.1343 | -85.7858
1.1346 | -85.8011
1.1349 | -85.8051
1.1353 | -85.8116
1.1356 | -85.8197
1.1359 | -85.831
1.1362 | -85.8391
1.1365 | -85.8448
1.1369 | -85.8503
1.1372 | -85.8567
1.1375 | -85.8616
1.1378 | -85.8689
1.1381 | -85.8804
1.1385 | -85.8885
1.1388 | -85.8922
1.1391 | -85.8952
1.1394 | -85.8937
1.1398 | -85.8924
1.1401 | -85.8844
1.1404 | -85.8931
1.1407 | -85.9078
1.141 | -85.9102
1.1414 | -85.9109
1.1417 | -85.9142
1.142 | -85.9173
1.1423 | -85.9157
1.1426 | -85.917
1.143 | -85.9158
1.1433 | -85.9102
1.1436 | -85.8948
1.1439 | -85.8788
1.1443 | -85.8783
1.1446 | -85.8778
1.1449 | -85.8691
1.1452 | -85.8599
1.1455 | -85.8504
1.1459 | -85.8433
1.1462 | -85.8336
1.1465 | -85.8201
1.1468 | -85.809
1.1471 | -85.7987
1.1475 | -85.7877
1.1478 | -85.7689
1.1481 | -85.7511
1.1484 | -85.7432
1.1487 | -85.7333
1.1491 | -85.7212
1.1494 | -85.7059
1.1497 | -85.6947
1.15 | -85.685
1.1504 | -85.6726
1.1507 | -85.6583
1.151 | -85.6467
1.1513 | -85.6395
1.1516 | -85.6276
1.152 | -85.6179
1.1523 | -85.6199
1.1526 | -85.606
1.1529 | -85.592
1.1532 | -85.5776
1.1536 | -85.553
1.1539 | -85.5351
1.1542 | -85.5215
1.1545 | -85.5179
1.1549 | -85.5071
1.1552 | -85.4942
1.1555 | -85.4757
1.1558 | -85.4585
1.1561 | -85.4448
1.1565 | -85.4287
1.1568 | -85.4161
1.1571 | -85.3985
1.1574 | -85.3767
1.1577 | -85.3608
1.1581 | -85.347
1.1584 | -85.3311
1.1587 | -85.3212
1.159 | -85.3123
1.1594 | -85.3031
1.1597 | -85.2896
1.16 | -85.2749
1.1603 | -85.2626
1.1606 | -85.2467
1.161 | -85.2331
1.1613 | -85.2176
1.1616 | -85.2035
1.1619 | -85.1903
1.1622 | -85.1775
1.1626 | -85.1654
1.1629 | -85.1542
1.1632 | -85.144
1.1635 | -85.1359
1.1639 | -85.1255
1.1642 | -85.1163
1.1645 | -85.1038
1.1648 | -85.0935
1.1651 | -85.0853
1.1655 | -85.0743
1.1658 | -85.07
1.1661 | -85.0622
1.1664 | -85.0539
1.1667 | -85.0445
1.1671 | -85.0357
1.1674 | -85.0258
1.1677 | -85.0168
1.168 | -85.0169
1.1683 | -85.0098
1.1687 | -85.0037
1.169 | -84.996
1.1693 | -84.9912
1.1696 | -84.9903
1.17 | -84.9869
1.1703 | -84.981
1.1706 | -84.9728
1.1709 | -84.9655
1.1712 | -84.9606
1.1716 | -84.9535
1.1719 | -84.9485
1.1722 | -84.9432
1.1725 | -84.9372
1.1728 | -84.9317
1.1732 | -84.9273
1.1735 | -84.9224
1.1738 | -84.9192
1.1741 | -84.9164
1.1745 | -84.9117
1.1748 | -84.9082
1.1751 | -84.9071
1.1754 | -84.9043
1.1757 | -84.9007
1.1761 | -84.896
1.1764 | -84.8889
1.1767 | -84.8866
1.177 | -84.8856
1.1773 | -84.8914
1.1777 | -84.8876
1.178 | -84.88
1.1783 | -84.8753
1.1786 | -84.8742
1.179 | -84.8733
1.1793 | -84.8747
1.1796 | -84.8749
1.1799 | -84.8772
1.1802 | -84.8736
1.1806 | -84.8725
1.1809 | -84.8694
1.1812 | -84.8635
1.1815 | -84.8623
1.1818 | -84.866
1.1822 | -84.8672
1.1825 | -84.8671
1.1828 | -84.8689
1.1831 | -84.8689
1.1835 | -84.8661
1.1838 | -84.866
1.1841 | -84.8669
1.1844 | -84.869
1.1847 | -84.8712
1.1851 | -84.8697
1.1854 | -84.8622
1.1857 | -84.8581
1.186 | -84.8579
1.1863 | -84.8595
1.1867 | -84.8592
1.187 | -84.8593
1.1873 | -84.8599
1.1876 | -84.8623
1.188 | -84.8658
1.1883 | -84.8628
1.1886 | -84.8586
1.1889 | -84.8579
1.1892 | -84.8563
1.1896 | -84.8596
1.1899 | -84.8646
1.1902 | -84.8653
1.1905 | -84.8641
1.1908 | -84.8632
1.1912 | -84.8656
1.1915 | -84.8677
1.1918 | -84.8681
1.1921 | -84.8676
1.1925 | -84.8673
1.1928 | -84.8703
1.1931 | -84.8713
1.1934 | -84.8705
1.1937 | -84.8692
1.1941 | -84.8681
1.1944 | -84.8689
1.1947 | -84.8717
1.195 | -84.8712
1.1953 | -84.8704
1.1957 | -84.8692
1.196 | -84.8698
1.1963 | -84.8698
1.1966 | -84.8717
1.1969 | -84.8674
1.1973 | -84.8599
1.1976 | -84.8542
1.1979 | -84.8423
1.1982 | -84.8326
1.1986 | -84.8293
1.1989 | -84.8255
1.1992 | -84.8222
1.1995 | -84.8195
1.1998 | -84.8067
1.2002 | -84.7967
1.2005 | -84.7928
1.2008 | -84.7886
1.2011 | -84.782
1.2014 | -84.7765
1.2018 | -84.7721
1.2021 | -84.7656
1.2024 | -84.7612
1.2027 | -84.7597
1.2031 | -84.7604
1.2034 | -84.756
1.2037 | -84.7519
1.204 | -84.7459
1.2043 | -84.7406
1.2047 | -84.7367
1.205 | -84.7306
1.2053 | -84.7272
1.2056 | -84.7232
1.2059 | -84.718
1.2063 | -84.7107
1.2066 | -84.7045
1.2069 | -84.6997
1.2072 | -84.6943
1.2076 | -84.6847
1.2079 | -84.6805
1.2082 | -84.6737
1.2085 | -84.6676
1.2088 | -84.6631
1.2092 | -84.6612
1.2095 | -84.6574
1.2098 | -84.6546
1.2101 | -84.6541
1.2104 | -84.6516
1.2108 | -84.645
1.2111 | -84.6397
1.2114 | -84.6364
1.2117 | -84.6296
1.2121 | -84.6223
1.2124 | -84.6132
1.2127 | -84.6092
1.213 | -84.6043
1.2133 | -84.598
1.2137 | -84.5905
1.214 | -84.5838
1.2143 | -84.5772
1.2146 | -84.5701
1.2149 | -84.5628
1.2153 | -84.5576
1.2156 | -84.5515
1.2159 | -84.5408
1.2162 | -84.5348
1.2166 | -84.5276
1.2169 | -84.5266
1.2172 | -84.5276
1.2175 | -84.5232
1.2178 | -84.5165
1.2182 | -84.5108
1.2185 | -84.5049
1.2188 | -84.4966
1.2191 | -84.4879
1.2194 | -84.4811
1.2198 | -84.4744
1.2201 | -84.469
1.2204 | -84.4625
1.2207 | -84.4576
1.221 | -84.4526
1.2214 | -84.4468
1.2217 | -84.4428
1.222 | -84.4386
1.2223 | -84.4314
1.2227 | -84.4256
1.223 | -84.4143
1.2233 | -84.4063
1.2236 | -84.4069
1.2239 | -84.4027
1.2243 | -84.3942
1.2246 | -84.3848
1.2249 | -84.3781
1.2252 | -84.3733
1.2255 | -84.3686
1.2259 | -84.3652
1.2262 | -84.3619
1.2265 | -84.3647
1.2268 | -84.3661
1.2272 | -84.3622
1.2275 | -84.3565
1.2278 | -84.3494
1.2281 | -84.3454
1.2284 | -84.3425
1.2288 | -84.3399
1.2291 | -84.3356
1.2294 | -84.3311
1.2297 | -84.3279
1.23 | -84.3228
1.2304 | -84.3162
1.2307 | -84.3168
1.231 | -84.3181
1.2313 | -84.3153
1.2317 | -84.3133
1.232 | -84.3118
1.2323 | -84.3094
1.2326 | -84.3086
1.2329 | -84.3126
1.2333 | -84.3124
1.2336 | -84.3086
1.2339 | -84.3058
1.2342 | -84.3065
1.2345 | -84.3073
1.2349 | -84.304
1.2352 | -84.2972
1.2355 | -84.2976
1.2358 | -84.2985
1.2362 | -84.3003
1.2365 | -84.301
1.2368 | -84.3032
1.2371 | -84.31
1.2374 | -84.3123
1.2378 | -84.3127
1.2381 | -84.3115
1.2384 | -84.3113
1.2387 | -84.3122
1.239 | -84.3165
1.2394 | -84.3197
1.2397 | -84.3217
1.24 | -84.3252
1.2403 | -84.3244
1.2406 | -84.321
1.241 | -84.3249
1.2413 | -84.3259
1.2416 | -84.3229
1.2419 | -84.3206
1.2423 | -84.3236
1.2426 | -84.3262
1.2429 | -84.326
1.2432 | -84.3268
1.2435 | -84.3288
1.2439 | -84.3317
1.2442 | -84.3333
1.2445 | -84.3356
1.2448 | -84.3388
1.2451 | -84.3349
1.2455 | -84.3331
1.2458 | -84.3307
1.2461 | -84.332
1.2464 | -84.334
1.2468 | -84.3335
1.2471 | -84.3344
1.2474 | -84.3341
1.2477 | -84.3367
1.248 | -84.3385
1.2484 | -84.3413
1.2487 | -84.3442
1.249 | -84.3467
1.2493 | -84.3481
1.2496 | -84.3507
1.25 | -84.3494
1.2503 | -84.3506
1.2506 | -84.3517
1.2509 | -84.3507
1.2513 | -84.3508
1.2516 | -84.3526
1.2519 | -84.3517
1.2522 | -84.3524
1.2525 | -84.3527
1.2529 | -84.354
1.2532 | -84.3549
1.2535 | -84.3549
1.2538 | -84.3565
1.2541 | -84.3562
1.2545 | -84.3552
1.2548 | -84.3572
1.2551 | -84.3591
1.2554 | -84.3627
1.2558 | -84.3641
1.2561 | -84.3618
1.2564 | -84.359
1.2567 | -84.3597
1.257 | -84.3616
1.2574 | -84.3623
1.2577 | -84.3613
1.258 | -84.3609
1.2583 | -84.3617
1.2586 | -84.3567
1.259 | -84.3532
1.2593 | -84.3549
1.2596 | -84.3553
1.2599 | -84.3534
1.2603 | -84.3517
1.2606 | -84.3505
1.2609 | -84.3512
1.2612 | -84.3518
1.2615 | -84.3547
1.2619 | -84.3545
1.2622 | -84.3499
1.2625 | -84.3474
1.2628 | -84.3463
1.2631 | -84.3485
1.2635 | -84.3533
1.2638 | -84.3526
1.2641 | -84.3478
1.2644 | -84.3476
1.2647 | -84.3436
1.2651 | -84.3408
1.2654 | -84.3372
1.2657 | -84.339
1.266 | -84.3371
1.2664 | -84.3373
1.2667 | -84.3403
1.267 | -84.3394
1.2673 | -84.3351
1.2676 | -84.3288
1.268 | -84.3259
1.2683 | -84.3243
1.2686 | -84.321
1.2689 | -84.3196
1.2692 | -84.3168
1.2696 | -84.3106
1.2699 | -84.3056
1.2702 | -84.3019
1.2705 | -84.2967
1.2709 | -84.2951
1.2712 | -84.2985
1.2715 | -84.2991
1.2718 | -84.2952
1.2721 | -84.2894
1.2725 | -84.2862
1.2728 | -84.2782
1.2731 | -84.2736
1.2734 | -84.2706
1.2737 | -84.2648
1.2741 | -84.2603
1.2744 | -84.2595
1.2747 | -84.2601
1.275 | -84.2587
1.2754 | -84.2553
1.2757 | -84.2486
1.276 | -84.24
1.2763 | -84.2343
1.2766 | -84.2307
1.277 | -84.2281
1.2773 | -84.2243
1.2776 | -84.2202
1.2779 | -84.2167
1.2782 | -84.2113
1.2786 | -84.207
1.2789 | -84.204
1.2792 | -84.2035
1.2795 | -84.2006
1.2799 | -84.1964
1.2802 | -84.1929
1.2805 | -84.1872
1.2808 | -84.1809
1.2811 | -84.175
1.2815 | -84.1719
1.2818 | -84.1689
1.2821 | -84.1679
1.2824 | -84.1647
1.2827 | -84.1619
1.2831 | -84.1576
1.2834 | -84.1534
1.2837 | -84.1534
1.284 | -84.1566
1.2844 | -84.1564
1.2847 | -84.1605
1.285 | -84.1637
1.2853 | -84.1624
1.2856 | -84.1618
1.286 | -84.1587
1.2863 | -84.1539
1.2866 | -84.157
1.2869 | -84.1535
1.2872 | -84.1417
1.2876 | -84.1324
1.2879 | -84.1315
1.2882 | -84.126
1.2885 | -84.1227
1.2888 | -84.1194
1.2892 | -84.1178
1.2895 | -84.1136
1.2898 | -84.1105
1.2901 | -84.1088
1.2905 | -84.1055
1.2908 | -84.1031
1.2911 | -84.1022
1.2914 | -84.0972
1.2917 | -84.0879
1.2921 | -84.0839
1.2924 | -84.0797
1.2927 | -84.0759
1.293 | -84.0738
1.2933 | -84.0723
1.2937 | -84.0709
1.294 | -84.0629
1.2943 | -84.0552
1.2946 | -84.0489
1.295 | -84.0485
1.2953 | -84.0432
1.2956 | -84.0447
1.2959 | -84.0434
1.2962 | -84.0407
1.2966 | -84.0366
1.2969 | -84.0321
1.2972 | -84.0295
1.2975 | -84.0256
1.2978 | -84.0214
1.2982 | -84.0167
1.2985 | -84.0123
1.2988 | -84.0088
1.2991 | -84.0065
1.2995 | -84.004
1.2998 | -84.0024
1.3001 | -84.0036
1.3004 | -84.0022
1.3007 | -83.9989
1.3011 | -83.9964
1.3014 | -83.9951
1.3017 | -83.9913
1.302 | -83.9856
1.3023 | -83.9823
1.3027 | -83.9802
1.303 | -83.9766
1.3033 | -83.9723
1.3036 | -83.9693
1.304 | -83.969
1.3043 | -83.967
1.3046 | -83.9664
1.3049 | -83.964
1.3052 | -83.9615
1.3056 | -83.9571
1.3059 | -83.9531
1.3062 | -83.9497
1.3065 | -83.946
1.3068 | -83.9435
1.3072 | -83.9376
1.3075 | -83.9346
1.3078 | -83.932
1.3081 | -83.929
1.3084 | -83.9246
1.3088 | -83.9224
1.3091 | -83.9177
1.3094 | -83.9143
1.3097 | -83.9126
1.3101 | -83.9065
1.3104 | -83.8965
1.3107 | -83.8906
1.311 | -83.8839
1.3113 | -83.8843
1.3117 | -83.8845
1.312 | -83.8785
1.3123 | -83.8699
1.3126 | -83.8611
1.3129 | -83.8554
1.3133 | -83.8533
1.3136 | -83.8516
1.3139 | -83.8487
1.3142 | -83.8413
1.3146 | -83.8289
1.3149 | -83.8201
1.3152 | -83.8106
1.3155 | -83.8047
1.3158 | -83.7969
1.3162 | -83.7886
1.3165 | -83.7793
1.3168 | -83.7684
1.3171 | -83.7504
1.3174 | -83.7362
1.3178 | -83.7271
1.3181 | -83.7215
1.3184 | -83.7129
1.3187 | -83.7007
1.3191 | -83.6926
1.3194 | -83.6847
1.3197 | -83.6758
1.32 | -83.6674
1.3203 | -83.6598
1.3207 | -83.6493
1.321 | -83.6404
1.3213 | -83.6314
1.3216 | -83.6189
1.3219 | -83.6085
1.3223 | -83.6
1.3226 | -83.591
1.3229 | -83.5848
1.3232 | -83.5761
1.3236 | -83.5665
1.3239 | -83.5559
1.3242 | -83.5485
1.3245 | -83.5443
1.3248 | -83.5353
1.3252 | -83.5259
1.3255 | -83.5051
1.3258 | -83.4935
1.3261 | -83.4827
1.3264 | -83.4718
1.3268 | -83.4631
1.3271 | -83.4542
1.3274 | -83.4453
1.3277 | -83.4358
1.3281 | -83.4257
1.3284 | -83.4137
1.3287 | -83.4024
1.329 | -83.3903
1.3293 | -83.3807
1.3297 | -83.3713
1.33 | -83.3631
1.3303 | -83.3577
1.3306 | -83.3481
1.3309 | -83.3388
1.3313 | -83.3328
1.3316 | -83.3281
1.3319 | -83.3203
1.3322 | -83.3094
1.3325 | -83.2995
1.3329 | -83.2907
1.3332 | -83.2807
1.3335 | -83.2716
1.3338 | -83.2628
1.3342 | -83.2546
1.3345 | -83.2448
1.3348 | -83.2383
1.3351 | -83.2302
1.3354 | -83.2256
1.3358 | -83.2211
1.3361 | -83.2154
1.3364 | -83.2073
1.3367 | -83.1993
1.337 | -83.1909
1.3374 | -83.1825
1.3377 | -83.1757
1.338 | -83.1667
1.3383 | -83.1578
1.3387 | -83.1469
1.339 | -83.1382
1.3393 | -83.1315
1.3396 | -83.1242
1.3399 | -83.1183
1.3403 | -83.1124
1.3406 | -83.1053
1.3409 | -83.098
1.3412 | -83.0916
1.3415 | -83.0857
1.3419 | -83.0782
1.3422 | -83.072
1.3425 | -83.0628
1.3428 | -83.0532
1.3432 | -83.0475
1.3435 | -83.0411
1.3438 | -83.0342
1.3441 | -83.0314
1.3444 | -83.0251
1.3448 | -83.0196
1.3451 | -83.0163
1.3454 | -83.0122
1.3457 | -83.0055
1.346 | -82.9997
1.3464 | -82.9959
1.3467 | -82.9927
1.347 | -82.9897
1.3473 | -82.9855
1.3477 | -82.9771
1.348 | -82.9716
1.3483 | -82.9686
1.3486 | -82.9645
1.3489 | -82.959
1.3493 | -82.955
1.3496 | -82.9538
1.3499 | -82.9524
1.3502 | -82.948
1.3505 | -82.9399
1.3509 | -82.9357
1.3512 | -82.9334
1.3515 | -82.9311
1.3518 | -82.9274
1.3521 | -82.9237
1.3525 | -82.9209
1.3528 | -82.9188
1.3531 | -82.9174
1.3534 | -82.9207
1.3538 | -82.9223
1.3541 | -82.922
1.3544 | -82.9182
1.3547 | -82.9146
1.355 | -82.914
1.3554 | -82.9152
1.3557 | -82.9137
1.356 | -82.9097
1.3563 | -82.9087
1.3566 | -82.9076
1.357 | -82.9091
1.3573 | -82.9086
1.3576 | -82.9094
1.3579 | -82.9101
1.3583 | -82.9114
1.3586 | -82.9131
1.3589 | -82.9161
1.3592 | -82.92
1.3595 | -82.9226
1.3599 | -82.9246
1.3602 | -82.9274
1.3605 | -82.9302
1.3608 | -82.9324
1.3611 | -82.9367
1.3615 | -82.9391
1.3618 | -82.9375
1.3621 | -82.9376
1.3624 | -82.9426
1.3628 | -82.9491
1.3631 | -82.954
1.3634 | -82.9579
1.3637 | -82.9617
1.364 | -82.9652
1.3644 | -82.971
1.3647 | -82.9755
1.365 | -82.9814
1.3653 | -82.986
1.3656 | -82.9908
1.366 | -82.9963
1.3663 | -83.0019
1.3666 | -83.0067
1.3669 | -83.0121
1.3673 | -83.0181
1.3676 | -83.0231
1.3679 | -83.0301
1.3682 | -83.0396
1.3685 | -83.0471
1.3689 | -83.0539
1.3692 | -83.0609
1.3695 | -83.0679
1.3698 | -83.0753
1.3701 | -83.0818
1.3705 | -83.0902
1.3708 | -83.0987
1.3711 | -83.1045
1.3714 | -83.1113
1.3718 | -83.1196
1.3721 | -83.1278
1.3724 | -83.1298
1.3727 | -83.1357
1.373 | -83.1429
1.3734 | -83.1493
1.3737 | -83.1582
1.374 | -83.1671
1.3743 | -83.1733
1.3746 | -83.1781
1.375 | -83.1852
1.3753 | -83.1935
1.3756 | -83.2016
1.3759 | -83.2114
1.3762 | -83.2209
1.3766 | -83.2271
1.3769 | -83.2326
1.3772 | -83.2425
1.3775 | -83.2518
1.3779 | -83.2602
1.3782 | -83.2682
1.3785 | -83.2755
1.3788 | -83.2847
1.3791 | -83.2952
1.3795 | -83.3056
1.3798 | -83.3179
1.3801 | -83.3316
1.3804 | -83.3451
1.3807 | -83.3537
1.3811 | -83.3652
1.3814 | -83.3742
1.3817 | -83.3852
1.382 | -83.3932
1.3824 | -83.4005
1.3827 | -83.4092
1.383 | -83.4183
1.3833 | -83.4265
1.3836 | -83.4351
1.384 | -83.4435
1.3843 | -83.451
1.3846 | -83.4592
1.3849 | -83.4711
1.3852 | -83.4841
1.3856 | -83.4939
1.3859 | -83.5035
1.3862 | -83.5133
1.3865 | -83.5202
1.3869 | -83.5236
1.3872 | -83.532
1.3875 | -83.544
1.3878 | -83.5533
1.3881 | -83.5647
1.3885 | -83.5737
1.3888 | -83.5835
1.3891 | -83.5895
1.3894 | -83.596
1.3897 | -83.6049
1.3901 | -83.6143
1.3904 | -83.6182
1.3907 | -83.6271
1.391 | -83.633
1.3914 | -83.6403
1.3917 | -83.6491
1.392 | -83.6563
1.3923 | -83.6634
1.3926 | -83.6706
1.393 | -83.6766
1.3933 | -83.6807
1.3936 | -83.6885
1.3939 | -83.6943
1.3942 | -83.7001
1.3946 | -83.7075
1.3949 | -83.7123
1.3952 | -83.7162
1.3955 | -83.724
1.3959 | -83.728
1.3962 | -83.7357
1.3965 | -83.7443
1.3968 | -83.7509
1.3971 | -83.7588
1.3975 | -83.7671
1.3978 | -83.7754
1.3981 | -83.7822
1.3984 | -83.7885
1.3987 | -83.7956
1.3991 | -83.8023
1.3994 | -83.8074
1.3997 | -83.8156
1.4 | -83.8221
1.4003 | -83.8279
1.4007 | -83.8324
1.401 | -83.8391
1.4013 | -83.8455
1.4016 | -83.8526
1.402 | -83.8625
1.4023 | -83.8697
1.4026 | -83.8776
1.4029 | -83.8853
1.4032 | -83.8922
1.4036 | -83.8949
1.4039 | -83.9012
1.4042 | -83.9064
1.4045 | -83.912
1.4048 | -83.9194
1.4052 | -83.9325
1.4055 | -83.9412
1.4058 | -83.9486
1.4061 | -83.9557
1.4065 | -83.9592
1.4068 | -83.966
1.4071 | -83.9763
1.4074 | -83.9875
1.4077 | -83.9973
1.4081 | -84.0027
1.4084 | -84.0064
1.4087 | -84.0076
1.409 | -84.0138
1.4093 | -84.023
1.4097 | -84.0349
1.41 | -84.0443
1.4103 | -84.0536
1.4106 | -84.0591
1.411 | -84.0666
1.4113 | -84.0764
1.4116 | -84.083
1.4119 | -84.0873
1.4122 | -84.1003
1.4126 | -84.1176
1.4129 | -84.1295
1.4132 | -84.1364
1.4135 | -84.1422
1.4138 | -84.1517
1.4142 | -84.1618
1.4145 | -84.1637
1.4148 | -84.1691
1.4151 | -84.1759
1.4155 | -84.1761
1.4158 | -84.176
1.4161 | -84.1817
1.4164 | -84.1862
1.4167 | -84.1915
1.4171 | -84.1998
1.4174 | -84.2059
1.4177 | -84.2102
1.418 | -84.2148
1.4183 | -84.2209
1.4187 | -84.2256
1.419 | -84.2311
1.4193 | -84.2362
1.4196 | -84.2419
1.42 | -84.2445
1.4203 | -84.2474
1.4206 | -84.2522
1.4209 | -84.2573
1.4212 | -84.2685
1.4216 | -84.2787
1.4219 | -84.2838
1.4222 | -84.2831
1.4225 | -84.2841
1.4228 | -84.2867
1.4232 | -84.2876
1.4235 | -84.2915
1.4238 | -84.2951
1.4241 | -84.2977
1.4244 | -84.3034
1.4248 | -84.3091
1.4251 | -84.3157
1.4254 | -84.3182
1.4257 | -84.3151
1.4261 | -84.3117
1.4264 | -84.3168
1.4267 | -84.32
1.427 | -84.3196
1.4273 | -84.3149
1.4277 | -84.3087
1.428 | -84.3072
1.4283 | -84.3063
1.4286 | -84.3087
1.4289 | -84.3063
1.4293 | -84.2982
1.4296 | -84.2978
1.4299 | -84.2986
1.4302 | -84.2995
1.4306 | -84.3009
1.4309 | -84.3028
1.4312 | -84.3034
1.4315 | -84.3041
1.4318 | -84.3056
1.4322 | -84.3064
1.4325 | -84.3053
1.4328 | -84.3042
1.4331 | -84.3045
1.4334 | -84.305
1.4338 | -84.3065
1.4341 | -84.3061
1.4344 | -84.3041
1.4347 | -84.304
1.4351 | -84.3034
1.4354 | -84.3055
1.4357 | -84.3068
1.436 | -84.3035
1.4363 | -84.3025
1.4367 | -84.2997
1.437 | -84.2955
1.4373 | -84.2928
1.4376 | -84.2887
1.4379 | -84.2861
1.4383 | -84.286
1.4386 | -84.283
1.4389 | -84.2792
1.4392 | -84.2739
1.4396 | -84.2784
1.4399 | -84.2791
1.4402 | -84.2747
1.4405 | -84.2712
1.4408 | -84.2682
1.4412 | -84.265
1.4415 | -84.2616
1.4418 | -84.2589
1.4421 | -84.2562
1.4424 | -84.2532
1.4428 | -84.2514
1.4431 | -84.2495
1.4434 | -84.246
1.4437 | -84.2378
1.4441 | -84.2315
1.4444 | -84.2274
1.4447 | -84.2261
1.445 | -84.2273
1.4453 | -84.2179
1.4457 | -84.2104
1.446 | -84.2049
1.4463 | -84.2008
1.4466 | -84.198
1.4469 | -84.1913
1.4473 | -84.1852
1.4476 | -84.1801
1.4479 | -84.1766
1.4482 | -84.178
1.4485 | -84.1752
1.4489 | -84.1704
1.4492 | -84.1652
1.4495 | -84.1574
1.4498 | -84.1473
1.4502 | -84.1391
1.4505 | -84.1336
1.4508 | -84.1265
1.4511 | -84.12
1.4514 | -84.1127
1.4518 | -84.1023
1.4521 | -84.0955
1.4524 | -84.092
1.4527 | -84.088
1.453 | -84.0806
1.4534 | -84.0709
1.4537 | -84.0595
1.454 | -84.0518
1.4543 | -84.0474
1.4547 | -84.0435
1.455 | -84.0355
1.4553 | -84.026
1.4556 | -84.0178
1.4559 | -84.0098
1.4563 | -84.0015
1.4566 | -83.9905
1.4569 | -83.981
1.4572 | -83.9725
1.4575 | -83.9646
1.4579 | -83.957
1.4582 | -83.9475
1.4585 | -83.9374
1.4588 | -83.9283
1.4592 | -83.922
1.4595 | -83.9153
1.4598 | -83.9108
1.4601 | -83.9078
1.4604 | -83.9034
1.4608 | -83.896
1.4611 | -83.8898
1.4614 | -83.8835
1.4617 | -83.8726
1.462 | -83.8618
1.4624 | -83.8549
1.4627 | -83.8495
1.463 | -83.8404
1.4633 | -83.8333
1.4637 | -83.8281
1.464 | -83.8222
1.4643 | -83.8142
1.4646 | -83.8059
1.4649 | -83.7943
1.4653 | -83.7851
1.4656 | -83.7762
1.4659 | -83.7697
1.4662 | -83.7615
1.4665 | -83.752
1.4669 | -83.7455
1.4672 | -83.737
1.4675 | -83.7267
1.4678 | -83.7165
1.4682 | -83.708
1.4685 | -83.6999
1.4688 | -83.6928
1.4691 | -83.6861
1.4694 | -83.6796
1.4698 | -83.6721
1.4701 | -83.6668
1.4704 | -83.6614
1.4707 | -83.6543
1.471 | -83.6443
1.4714 | -83.6359
1.4717 | -83.6332
1.472 | -83.6304
1.4723 | -83.6235
1.4726 | -83.6183
1.473 | -83.6127
1.4733 | -83.6065
1.4736 | -83.6022
1.4739 | -83.5986
1.4743 | -83.592
1.4746 | -83.5854
1.4749 | -83.5793
1.4752 | -83.5731
1.4755 | -83.5674
1.4759 | -83.5603
1.4762 | -83.5555
1.4765 | -83.5534
1.4768 | -83.5475
1.4771 | -83.5418
1.4775 | -83.5321
1.4778 | -83.5185
1.4781 | -83.5132
1.4784 | -83.5074
1.4788 | -83.5017
1.4791 | -83.4952
1.4794 | -83.4909
1.4797 | -83.4876
1.48 | -83.4827
1.4804 | -83.4773
1.4807 | -83.4721
1.481 | -83.4675
1.4813 | -83.4592
1.4816 | -83.4463
1.482 | -83.4363
1.4823 | -83.4323
1.4826 | -83.4213
1.4829 | -83.4124
1.4833 | -83.4091
1.4836 | -83.4014
1.4839 | -83.3959
1.4842 | -83.3891
1.4845 | -83.3782
1.4849 | -83.3697
1.4852 | -83.3631
1.4855 | -83.3578
1.4858 | -83.3532
1.4861 | -83.3457
1.4865 | -83.3347
1.4868 | -83.3273
1.4871 | -83.3257
1.4874 | -83.3204
1.4878 | -83.3143
1.4881 | -83.3066
1.4884 | -83.2994
1.4887 | -83.2939
1.489 | -83.2883
1.4894 | -83.2847
1.4897 | -83.2794
1.49 | -83.2718
1.4903 | -83.266
1.4906 | -83.2606
1.491 | -83.2556
1.4913 | -83.2494
1.4916 | -83.2403
1.4919 | -83.2344
1.4923 | -83.2288
1.4926 | -83.2198
1.4929 | -83.2139
1.4932 | -83.2097
1.4935 | -83.2052
1.4939 | -83.2006
1.4942 | -83.1974
1.4945 | -83.1915
1.4948 | -83.1862
1.4951 | -83.1807
1.4955 | -83.1743
1.4958 | -83.17
1.4961 | -83.162
1.4964 | -83.1511
1.4967 | -83.147
1.4971 | -83.1428
1.4974 | -83.1365
1.4977 | -83.1302
1.498 | -83.1237
1.4984 | -83.117
1.4987 | -83.1107
1.499 | -83.1046
1.4993 | -83.0979
1.4996 | -83.0889
1.5 | -83.0824
1.5003 | -83.078
1.5006 | -83.072
1.5009 | -83.0665
1.5012 | -83.0597
1.5016 | -83.0528
1.5019 | -83.0313
1.5022 | -83.0258
1.5025 | -83.0181
1.5029 | -83.0082
1.5032 | -82.9991
1.5035 | -82.9906
1.5038 | -82.9829
1.5041 | -82.976
1.5045 | -82.9678
1.5048 | -82.9598
1.5051 | -82.9549
1.5054 | -82.9448
1.5057 | -82.9381
1.5061 | -82.9288
1.5064 | -82.9233
1.5067 | -82.9161
1.507 | -82.9101
1.5074 | -82.9012
1.5077 | -82.8829
1.508 | -82.8717
1.5083 | -82.8583
1.5086 | -82.8441
1.509 | -82.8325
1.5093 | -82.8233
1.5096 | -82.815
1.5099 | -82.8068
1.5102 | -82.8011
1.5106 | -82.7923
1.5109 | -82.7658
1.5112 | -82.7453
1.5115 | -82.7283
1.5119 | -82.7162
1.5122 | -82.7041
1.5125 | -82.6932
1.5128 | -82.6833
1.5131 | -82.6731
1.5135 | -82.6605
1.5138 | -82.6474
1.5141 | -82.6357
1.5144 | -82.6232
1.5147 | -82.613
1.5151 | -82.6013
1.5154 | -82.5891
1.5157 | -82.5751
1.516 | -82.5604
1.5163 | -82.5474
1.5167 | -82.5386
1.517 | -82.5263
1.5173 | -82.5082
1.5176 | -82.4968
1.518 | -82.4876
1.5183 | -82.4749
1.5186 | -82.4633
1.5189 | -82.4508
1.5192 | -82.4382
1.5196 | -82.4264
1.5199 | -82.4151
1.5202 | -82.401
1.5205 | -82.3882
1.5208 | -82.3768
1.5212 | -82.3652
1.5215 | -82.3531
1.5218 | -82.3414
1.5221 | -82.3297
1.5225 | -82.3185
1.5228 | -82.3089
1.5231 | -82.2983
1.5234 | -82.2848
1.5237 | -82.2732
1.5241 | -82.2621
1.5244 | -82.2496
1.5247 | -82.2389
1.525 | -82.2314
1.5253 | -82.2222
1.5257 | -82.2101
1.526 | -82.1987
1.5263 | -82.1862
1.5266 | -82.176
1.527 | -82.1661
1.5273 | -82.154
1.5276 | -82.139
1.5279 | -82.1251
1.5282 | -82.1132
1.5286 | -82.1053
1.5289 | -82.0978
1.5292 | -82.0847
1.5295 | -82.0727
1.5298 | -82.0626
1.5302 | -82.0528
1.5305 | -82.0444
1.5308 | -82.0351
1.5311 | -82.0252
1.5315 | -82.0152
1.5318 | -82.0046
1.5321 | -81.9925
1.5324 | -81.982
1.5327 | -81.976
1.5331 | -81.9659
1.5334 | -81.9509
1.5337 | -81.9373
1.534 | -81.9258
1.5343 | -81.9113
1.5347 | -81.8925
1.535 | -81.8763
1.5353 | -81.8634
1.5356 | -81.8507
1.5359 | -81.8395
1.5363 | -81.8301
1.5366 | -81.8191
1.5369 | -81.8077
1.5372 | -81.7973
1.5376 | -81.7876
1.5379 | -81.7765
1.5382 | -81.7611
1.5385 | -81.7514
1.5388 | -81.7436
1.5392 | -81.7343
1.5395 | -81.7246
1.5398 | -81.7148
1.5401 | -81.7048
1.5404 | -81.6945
1.5408 | -81.6848
1.5411 | -81.676
1.5414 | -81.6672
1.5417 | -81.6593
1.5421 | -81.6506
1.5424 | -81.6415
1.5427 | -81.6338
1.543 | -81.6268
1.5433 | -81.6207
1.5437 | -81.6129
1.544 | -81.6044
1.5443 | -81.5907
1.5446 | -81.5794
1.5449 | -81.5722
1.5453 | -81.5648
1.5456 | -81.5575
1.5459 | -81.548
1.5462 | -81.5386
1.5466 | -81.5296
1.5469 | -81.5208
1.5472 | -81.5133
1.5475 | -81.507
1.5478 | -81.5025
1.5482 | -81.5019
1.5485 | -81.4967
1.5488 | -81.4906
1.5491 | -81.4854
1.5494 | -81.4803
1.5498 | -81.478
1.5501 | -81.473
1.5504 | -81.4653
1.5507 | -81.4584
1.5511 | -81.4541
1.5514 | -81.4499
1.5517 | -81.4464
1.552 | -81.4407
1.5523 | -81.4325
1.5527 | -81.4272
1.553 | -81.4215
1.5533 | -81.4127
1.5536 | -81.4069
1.5539 | -81.4017
1.5543 | -81.3982
1.5546 | -81.3961
1.5549 | -81.3931
1.5552 | -81.3901
1.5556 | -81.3873
1.5559 | -81.3828
1.5562 | -81.3775
1.5565 | -81.3739
1.5568 | -81.371
1.5572 | -81.3689
1.5575 | -81.3724
1.5578 | -81.3707
1.5581 | -81.3681
1.5584 | -81.3671
1.5588 | -81.3701
1.5591 | -81.371
1.5594 | -81.3697
1.5597 | -81.3679
1.56 | -81.3665
1.5604 | -81.3659
1.5607 | -81.3657
1.561 | -81.3655
1.5613 | -81.3635
1.5617 | -81.3575
1.562 | -81.3568
1.5623 | -81.3564
1.5626 | -81.356
1.5629 | -81.3568
1.5633 | -81.3574
1.5636 | -81.3587
1.5639 | -81.3589
1.5642 | -81.3581
1.5645 | -81.3585
1.5649 | -81.3574
1.5652 | -81.3557
1.5655 | -81.3567
1.5658 | -81.3572
1.5662 | -81.3582
1.5665 | -81.3596
1.5668 | -81.3667
1.5671 | -81.3686
1.5674 | -81.3707
1.5678 | -81.3731
1.5681 | -81.3748
1.5684 | -81.3769
1.5687 | -81.38
1.569 | -81.3825
1.5694 | -81.3848
1.5697 | -81.3875
1.57 | -81.39
1.5703 | -81.3921
1.5707 | -81.3963
1.571 | -81.4024
1.5713 | -81.4071
1.5716 | -81.4121
1.5719 | -81.4138
1.5723 | -81.4121
1.5726 | -81.4139
1.5729 | -81.4166
1.5732 | -81.4175
1.5735 | -81.4171
1.5739 | -81.4194
1.5742 | -81.4226
1.5745 | -81.4237
1.5748 | -81.4239
1.5752 | -81.427
1.5755 | -81.4307
1.5758 | -81.4382
1.5761 | -81.4416
1.5764 | -81.4423
1.5768 | -81.4427
1.5771 | -81.4466
1.5774 | -81.453
1.5777 | -81.4599
1.578 | -81.4645
1.5784 | -81.4683
1.5787 | -81.4728
1.579 | -81.4773
1.5793 | -81.4821
1.5797 | -81.4868
1.58 | -81.4903
1.5803 | -81.4934
1.5806 | -81.4971
1.5809 | -81.5019
1.5813 | -81.5046
1.5816 | -81.5047
1.5819 | -81.5057
1.5822 | -81.5089
1.5825 | -81.5138
1.5829 | -81.5164
1.5832 | -81.5165
1.5835 | -81.5175
1.5838 | -81.5198
1.5841 | -81.5263
1.5845 | -81.5355
1.5848 | -81.5399
1.5851 | -81.5403
1.5854 | -81.5345
1.5858 | -81.5311
1.5861 | -81.5332
1.5864 | -81.5362
1.5867 | -81.5393
1.587 | -81.5421
1.5874 | -81.5441
1.5877 | -81.5444
1.588 | -81.5449
1.5883 | -81.5478
1.5886 | -81.5531
1.589 | -81.5582
1.5893 | -81.5591
1.5896 | -81.5604
1.5899 | -81.5602
1.5903 | -81.5587
1.5906 | -81.5588
1.5909 | -81.5604
1.5912 | -81.5611
1.5915 | -81.5637
1.5919 | -81.5688
1.5922 | -81.5717
1.5925 | -81.5741
1.5928 | -81.5757
1.5931 | -81.5762
1.5935 | -81.5789
1.5938 | -81.5838
1.5941 | -81.5852
1.5944 | -81.5849
1.5948 | -81.5849
1.5951 | -81.5852
1.5954 | -81.5861
1.5957 | -81.5897
1.596 | -81.59
1.5964 | -81.586
1.5967 | -81.585
1.597 | -81.5838
1.5973 | -81.5832
1.5976 | -81.5815
1.598 | -81.5778
1.5983 | -81.577
1.5986 | -81.5758
1.5989 | -81.5763
1.5993 | -81.5786
1.5996 | -81.5784
1.5999 | -81.5775
1.6002 | -81.5778
1.6005 | -81.5752
1.6009 | -81.5718
1.6012 | -81.5683
1.6015 | -81.5659
1.6018 | -81.5642
1.6021 | -81.5626
1.6025 | -81.5611
1.6028 | -81.5597
1.6031 | -81.5584
1.6034 | -81.5565
1.6038 | -81.5532
1.6041 | -81.5518
1.6044 | -81.5514
1.6047 | -81.5503
1.605 | -81.5468
1.6054 | -81.543
1.6057 | -81.54
1.606 | -81.5364
1.6063 | -81.5321
1.6066 | -81.5288
1.607 | -81.5233
1.6073 | -81.5202
1.6076 | -81.5178
1.6079 | -81.515
1.6082 | -81.5097
1.6086 | -81.502
1.6089 | -81.4967
1.6092 | -81.4936
1.6095 | -81.4919
1.6099 | -81.4897
1.6102 | -81.4897
1.6105 | -81.4921
1.6108 | -81.4892
1.6111 | -81.4871
1.6115 | -81.486
1.6118 | -81.4838
1.6121 | -81.4804
1.6124 | -81.4778
1.6127 | -81.4729
1.6131 | -81.4696
1.6134 | -81.4653
1.6137 | -81.4616
1.614 | -81.4583
1.6144 | -81.4567
1.6147 | -81.4519
1.615 | -81.4477
1.6153 | -81.4438
1.6156 | -81.4399
1.616 | -81.4339
1.6163 | -81.4289
1.6166 | -81.4297
1.6169 | -81.4286
1.6172 | -81.428
1.6176 | -81.4273
1.6179 | -81.4243
1.6182 | -81.4173
1.6185 | -81.4121
1.6189 | -81.4072
1.6192 | -81.4004
1.6195 | -81.3968
1.6198 | -81.3974
1.6201 | -81.3994
1.6205 | -81.4005
1.6208 | -81.3999
1.6211 | -81.397
1.6214 | -81.3939
1.6217 | -81.39
1.6221 | -81.3853
1.6224 | -81.3828
1.6227 | -81.377
1.623 | -81.3736
1.6234 | -81.372
1.6237 | -81.3706
1.624 | -81.3678
1.6243 | -81.3617
1.6246 | -81.3571
1.625 | -81.3545
1.6253 | -81.3533
1.6256 | -81.3528
1.6259 | -81.3511
1.6262 | -81.3481
1.6266 | -81.3474
1.6269 | -81.3481
1.6272 | -81.3458
1.6275 | -81.3408
1.6279 | -81.337
1.6282 | -81.3343
1.6285 | -81.3292
1.6288 | -81.3211
1.6291 | -81.3143
1.6295 | -81.3089
1.6298 | -81.3049
1.6301 | -81.3033
1.6304 | -81.3021
1.6307 | -81.2994
1.6311 | -81.295
1.6314 | -81.2925
1.6317 | -81.2958
1.632 | -81.2914
1.6323 | -81.2878
1.6327 | -81.2855
1.633 | -81.2831
1.6333 | -81.2819
1.6336 | -81.2815
1.634 | -81.2801
1.6343 | -81.2788
1.6346 | -81.2793
1.6349 | -81.2781
1.6352 | -81.2766
1.6356 | -81.276
1.6359 | -81.2767
1.6362 | -81.2773
1.6365 | -81.2743
1.6368 | -81.2722
1.6372 | -81.2725
1.6375 | -81.2715
1.6378 | -81.2728
1.6381 | -81.2733
1.6385 | -81.2707
1.6388 | -81.2665
1.6391 | -81.265
1.6394 | -81.2665
1.6397 | -81.2662
1.6401 | -81.2666
1.6404 | -81.268
1.6407 | -81.2679
1.641 | -81.2649
1.6413 | -81.2658
1.6417 | -81.2666
1.642 | -81.2657
1.6423 | -81.2674
1.6426 | -81.2678
1.643 | -81.2696
1.6433 | -81.2722
1.6436 | -81.2741
1.6439 | -81.2764
1.6442 | -81.2806
1.6446 | -81.2828
1.6449 | -81.2837
1.6452 | -81.2856
1.6455 | -81.2892
1.6458 | -81.2921
1.6462 | -81.2939
1.6465 | -81.2973
1.6468 | -81.2989
1.6471 | -81.3006
1.6475 | -81.303
1.6478 | -81.3048
1.6481 | -81.3072
1.6484 | -81.3091
1.6487 | -81.3125
1.6491 | -81.3193
1.6494 | -81.3221
1.6497 | -81.3243
1.65 | -81.3296
1.6503 | -81.3327
1.6507 | -81.3352
1.651 | -81.333
1.6513 | -81.3322
1.6516 | -81.3354
1.652 | -81.3389
1.6523 | -81.3423
1.6526 | -81.3454
1.6529 | -81.349
1.6532 | -81.3525
1.6536 | -81.3551
1.6539 | -81.3562
1.6542 | -81.3568
1.6545 | -81.3565
1.6548 | -81.3581
1.6552 | -81.3611
1.6555 | -81.3646
1.6558 | -81.3722
1.6561 | -81.3768
1.6564 | -81.3788
1.6568 | -81.3811
1.6571 | -81.384
1.6574 | -81.3864
1.6577 | -81.3883
1.6581 | -81.3902
1.6584 | -81.393
1.6587 | -81.3957
1.659 | -81.3982
1.6593 | -81.4024
1.6597 | -81.4042
1.66 | -81.4066
1.6603 | -81.4096
1.6606 | -81.4128
1.6609 | -81.4159
1.6613 | -81.4177
1.6616 | -81.4184
1.6619 | -81.4193
1.6622 | -81.4216
1.6626 | -81.4249
1.6629 | -81.428
1.6632 | -81.4272
1.6635 | -81.427
1.6638 | -81.4287
1.6642 | -81.433
1.6645 | -81.4393
1.6648 | -81.4418
1.6651 | -81.4442
1.6654 | -81.4491
1.6658 | -81.4523
1.6661 | -81.4556
1.6664 | -81.459
1.6667 | -81.4607
1.6671 | -81.4635
1.6674 | -81.4642
1.6677 | -81.464
1.668 | -81.4661
1.6683 | -81.4694
1.6687 | -81.465
1.669 | -81.4658
1.6693 | -81.466
1.6696 | -81.4667
1.6699 | -81.4678
1.6703 | -81.4693
1.6706 | -81.4703
1.6709 | -81.4698
1.6712 | -81.4704
1.6716 | -81.4699
1.6719 | -81.4705
1.6722 | -81.4712
1.6725 | -81.4722
1.6728 | -81.4732
1.6732 | -81.4746
1.6735 | -81.476
1.6738 | -81.4765
1.6741 | -81.4795
1.6744 | -81.4844
1.6748 | -81.4815
1.6751 | -81.4799
1.6754 | -81.48
1.6757 | -81.48
1.676 | -81.4799
1.6764 | -81.4798
1.6767 | -81.4819
1.677 | -81.4827
1.6773 | -81.4801
1.6777 | -81.4774
1.678 | -81.4659
1.6783 | -81.4679
1.6786 | -81.4681
1.6789 | -81.4689
1.6793 | -81.4683
1.6796 | -81.4648
1.6799 | -81.4619
1.6802 | -81.4605
1.6805 | -81.4597
1.6809 | -81.4605
1.6812 | -81.4612
1.6815 | -81.4606
1.6818 | -81.4594
1.6822 | -81.458
1.6825 | -81.4575
1.6828 | -81.4565
1.6831 | -81.4554
1.6834 | -81.4545
1.6838 | -81.4546
1.6841 | -81.4547
1.6844 | -81.4533
1.6847 | -81.4519
1.685 | -81.4497
1.6854 | -81.4505
1.6857 | -81.4551
1.686 | -81.4577
1.6863 | -81.4576
1.6867 | -81.4598
1.687 | -81.4631
1.6873 | -81.4614
1.6876 | -81.4625
1.6879 | -81.4601
1.6883 | -81.4598
1.6886 | -81.4595
1.6889 | -81.4567
1.6892 | -81.4538
1.6895 | -81.4509
1.6899 | -81.4494
1.6902 | -81.4466
1.6905 | -81.4463
1.6908 | -81.4474
1.6912 | -81.4455
1.6915 | -81.4435
1.6918 | -81.4419
1.6921 | -81.4401
1.6924 | -81.4384
1.6928 | -81.4382
1.6931 | -81.4376
1.6934 | -81.4361
1.6937 | -81.4336
1.694 | -81.4297
1.6944 | -81.429
1.6947 | -81.4288
1.695 | -81.4265
1.6953 | -81.4232
1.6957 | -81.4233
1.696 | -81.4253
1.6963 | -81.4268
1.6966 | -81.4257
1.6969 | -81.4246
1.6973 | -81.4251
1.6976 | -81.4252
1.6979 | -81.425
1.6982 | -81.424
1.6985 | -81.4221
1.6989 | -81.4205
1.6992 | -81.4195
1.6995 | -81.4184
1.6998 | -81.4201
1.7001 | -81.4282
1.7005 | -81.4332
1.7008 | -81.4335
1.7011 | -81.4326
1.7014 | -81.4308
1.7018 | -81.4304
1.7021 | -81.431
1.7024 | -81.4324
1.7027 | -81.4338
1.703 | -81.435
1.7034 | -81.4365
1.7037 | -81.4356
1.704 | -81.4362
1.7043 | -81.4399
1.7046 | -81.4418
1.705 | -81.4422
1.7053 | -81.444
1.7056 | -81.4452
1.7059 | -81.4464
1.7063 | -81.4469
1.7066 | -81.4467
1.7069 | -81.4465
1.7072 | -81.4467
1.7075 | -81.4463
1.7079 | -81.4467
1.7082 | -81.4473
1.7085 | -81.4472
1.7088 | -81.4471
1.7091 | -81.4487
1.7095 | -81.4489
1.7098 | -81.4498
1.7101 | -81.4508
1.7104 | -81.4524
1.7108 | -81.4531
1.7111 | -81.4528
1.7114 | -81.4551
1.7117 | -81.4596
1.712 | -81.4614
1.7124 | -81.4599
1.7127 | -81.4624
1.713 | -81.4643
1.7133 | -81.4669
1.7136 | -81.4695
1.714 | -81.4725
1.7143 | -81.4743
1.7146 | -81.4755
1.7149 | -81.4802
1.7153 | -81.4809
1.7156 | -81.48
1.7159 | -81.4789
1.7162 | -81.4782
1.7165 | -81.4805
1.7169 | -81.4836
1.7172 | -81.4858
1.7175 | -81.4849
1.7178 | -81.4865
1.7181 | -81.4886
1.7185 | -81.4907
1.7188 | -81.4932
1.7191 | -81.4951
1.7194 | -81.4947
1.7197 | -81.4982
1.7201 | -81.5007
1.7204 | -81.5044
1.7207 | -81.5092
1.721 | -81.5145
1.7214 | -81.5155
1.7217 | -81.518
1.722 | -81.5223
1.7223 | -81.5274
1.7226 | -81.5315
1.723 | -81.5369
1.7233 | -81.5409
1.7236 | -81.5428
1.7239 | -81.5461
1.7242 | -81.5523
1.7246 | -81.5587
1.7249 | -81.5659
1.7252 | -81.5697
1.7255 | -81.5719
1.7259 | -81.5754
1.7262 | -81.5793
1.7265 | -81.5809
1.7268 | -81.5837
1.7271 | -81.5875
1.7275 | -81.5912
1.7278 | -81.5953
1.7281 | -81.5988
1.7284 | -81.6026
1.7287 | -81.6068
1.7291 | -81.6124
1.7294 | -81.6175
1.7297 | -81.6218
1.73 | -81.6278
1.7304 | -81.6336
1.7307 | -81.6386
1.731 | -81.6438
1.7313 | -81.6498
1.7316 | -81.6547
1.732 | -81.6588
1.7323 | -81.6643
1.7326 | -81.6717
1.7329 | -81.6798
1.7332 | -81.6845
1.7336 | -81.6871
1.7339 | -81.6923
1.7342 | -81.6972
1.7345 | -81.7022
1.7349 | -81.7078
1.7352 | -81.7168
1.7355 | -81.7238
1.7358 | -81.7304
1.7361 | -81.7378
1.7365 | -81.7429
1.7368 | -81.7484
1.7371 | -81.7548
1.7374 | -81.7617
1.7377 | -81.7679
1.7381 | -81.7739
1.7384 | -81.7797
1.7387 | -81.7859
1.739 | -81.7907
1.7394 | -81.7945
1.7397 | -81.8018
1.74 | -81.8099
1.7403 | -81.8142
1.7406 | -81.819
1.741 | -81.8244
1.7413 | -81.8284
1.7416 | -81.8293
1.7419 | -81.8315
1.7422 | -81.8363
1.7426 | -81.8421
1.7429 | -81.8498
1.7432 | -81.8556
1.7435 | -81.8597
1.7438 | -81.8599
1.7442 | -81.8593
1.7445 | -81.8645
1.7448 | -81.8703
1.7451 | -81.8775
1.7455 | -81.8851
1.7458 | -81.8908
1.7461 | -81.8998
1.7464 | -81.9111
1.7467 | -81.9172
1.7471 | -81.9215
1.7474 | -81.9258
1.7477 | -81.9299
1.748 | -81.9331
1.7483 | -81.9383
1.7487 | -81.945
1.749 | -81.9507
1.7493 | -81.9558
1.7496 | -81.9614
1.75 | -81.9699
1.7503 | -81.9773
1.7506 | -81.9835
1.7509 | -81.9897
1.7512 | -81.9942
1.7516 | -81.9988
1.7519 | -82.0028
1.7522 | -82.0068
1.7525 | -82.0096
1.7528 | -82.012
1.7532 | -82.0138
1.7535 | -82.0163
1.7538 | -82.0215
1.7541 | -82.0255
1.7545 | -82.0302
1.7548 | -82.0391
1.7551 | -82.0456
1.7554 | -82.0493
1.7557 | -82.0551
1.7561 | -82.0618
1.7564 | -82.0651
1.7567 | -82.0684
1.757 | -82.077
1.7573 | -82.0828
1.7577 | -82.0857
1.758 | -82.0875
1.7583 | -82.0897
1.7586 | -82.0907
1.759 | -82.0957
1.7593 | -82.1025
1.7596 | -82.1083
1.7599 | -82.1119
1.7602 | -82.1155
1.7606 | -82.1205
1.7609 | -82.1224
1.7612 | -82.1227
1.7615 | -82.1253
1.7618 | -82.1265
1.7622 | -82.1291
1.7625 | -82.1299
1.7628 | -82.1277
1.7631 | -82.1273
1.7635 | -82.13
1.7638 | -82.1335
1.7641 | -82.1387
1.7644 | -82.1448
1.7647 | -82.1461
1.7651 | -82.1425
1.7654 | -82.1392
1.7657 | -82.1418
1.766 | -82.1427
1.7663 | -82.1437
1.7667 | -82.1458
1.767 | -82.1461
1.7673 | -82.1471
1.7676 | -82.1486
1.7679 | -82.145
1.7683 | -82.1355
1.7686 | -82.1295
1.7689 | -82.1232
1.7692 | -82.1188
1.7696 | -82.116
1.7699 | -82.1139
1.7702 | -82.1158
1.7705 | -82.1032
1.7708 | -82.1062
1.7712 | -82.1033
1.7715 | -82.1004
1.7718 | -82.0984
1.7721 | -82.096
1.7724 | -82.0941
1.7728 | -82.0904
1.7731 | -82.088
1.7734 | -82.088
1.7737 | -82.0811
1.7741 | -82.077
1.7744 | -82.0741
1.7747 | -82.0674
1.775 | -82.0584
1.7753 | -82.0524
1.7757 | -82.0499
1.776 | -82.0471
1.7763 | -82.0431
1.7766 | -82.0396
1.7769 | -82.0384
1.7773 | -82.0373
1.7776 | -82.0351
1.7779 | -82.0302
1.7782 | -82.0217
1.7786 | -82.0097
1.7789 | -81.9995
1.7792 | -81.9919
1.7795 | -81.9866
1.7798 | -81.9788
1.7802 | -81.9731
1.7805 | -81.9649
1.7808 | -81.9584
1.7811 | -81.9526
1.7814 | -81.9481
1.7818 | -81.9442
1.7821 | -81.94
1.7824 | -81.9348
1.7827 | -81.9304
1.7831 | -81.9252
1.7834 | -81.9209
1.7837 | -81.9165
1.784 | -81.9108
1.7843 | -81.9038
1.7847 | -81.8941
1.785 | -81.8867
1.7853 | -81.879
1.7856 | -81.8716
1.7859 | -81.8643
1.7863 | -81.8567
1.7866 | -81.8483
1.7869 | -81.8407
1.7872 | -81.8308
1.7875 | -81.8241
1.7879 | -81.8181
1.7882 | -81.8112
1.7885 | -81.8022
1.7888 | -81.7931
1.7892 | -81.782
1.7895 | -81.7738
1.7898 | -81.7661
1.7901 | -81.7574
1.7904 | -81.7475
1.7908 | -81.7376
1.7911 | -81.7285
1.7914 | -81.7159
1.7917 | -81.7053
1.792 | -81.6972
1.7924 | -81.6877
1.7927 | -81.6762
1.793 | -81.6663
1.7933 | -81.6566
1.7937 | -81.6477
1.794 | -81.637
1.7943 | -81.6278
1.7946 | -81.619
1.7949 | -81.6095
1.7953 | -81.6016
1.7956 | -81.5964
1.7959 | -81.5865
1.7962 | -81.5762
1.7965 | -81.5659
1.7969 | -81.5558
1.7972 | -81.5459
1.7975 | -81.5351
1.7978 | -81.5271
1.7982 | -81.5203
1.7985 | -81.5094
1.7988 | -81.5039
1.7991 | -81.5007
1.7994 | -81.4922
1.7998 | -81.4809
1.8001 | -81.4697
1.8004 | -81.4541
1.8007 | -81.4386
1.801 | -81.4274
1.8014 | -81.4174
1.8017 | -81.4079
1.802 | -81.3976
1.8023 | -81.3841
1.8027 | -81.3689
1.803 | -81.3574
1.8033 | -81.3476
1.8036 | -81.3383
1.8039 | -81.3275
1.8043 | -81.3134
1.8046 | -81.3017
1.8049 | -81.2917
1.8052 | -81.2822
1.8055 | -81.2718
1.8059 | -81.262
1.8062 | -81.2526
1.8065 | -81.243
1.8068 | -81.2307
1.8072 | -81.2155
1.8075 | -81.1985
1.8078 | -81.1897
1.8081 | -81.18
1.8084 | -81.1702
1.8088 | -81.1614
1.8091 | -81.1527
1.8094 | -81.1439
1.8097 | -81.135
1.81 | -81.1272
1.8104 | -81.1172
1.8107 | -81.1067
1.811 | -81.0966
1.8113 | -81.0887
1.8116 | -81.0807
1.812 | -81.0715
1.8123 | -81.064
1.8126 | -81.056
1.8129 | -81.047
1.8133 | -81.0385
1.8136 | -81.0305
1.8139 | -81.0197
1.8142 | -81.0096
1.8145 | -80.9994
1.8149 | -80.9905
1.8152 | -80.9811
1.8155 | -80.9703
1.8158 | -80.9623
1.8161 | -80.955
1.8165 | -80.9455
1.8168 | -80.9379
1.8171 | -80.9344
1.8174 | -80.9287
1.8178 | -80.9242
1.8181 | -80.9187
1.8184 | -80.9104
1.8187 | -80.9016
1.819 | -80.8945
1.8194 | -80.8882
1.8197 | -80.8807
1.82 | -80.8712
1.8203 | -80.8637
1.8206 | -80.8572
1.821 | -80.8514
1.8213 | -80.8441
1.8216 | -80.8366
1.8219 | -80.8291
1.8223 | -80.8225
1.8226 | -80.8161
1.8229 | -80.808
1.8232 | -80.8019
1.8235 | -80.796
1.8239 | -80.7903
1.8242 | -80.7853
1.8245 | -80.7769
1.8248 | -80.7656
1.8251 | -80.759
1.8255 | -80.754
1.8258 | -80.7489
1.8261 | -80.7487
1.8264 | -80.7434
1.8268 | -80.7392
1.8271 | -80.7345
1.8274 | -80.7299
1.8277 | -80.7267
1.828 | -80.724
1.8284 | -80.7199
1.8287 | -80.7171
1.829 | -80.7137
1.8293 | -80.7096
1.8296 | -80.7058
1.83 | -80.7021
1.8303 | -80.6984
1.8306 | -80.6954
1.8309 | -80.6915
1.8313 | -80.6871
1.8316 | -80.6844
1.8319 | -80.6825
1.8322 | -80.6804
1.8325 | -80.6771
1.8329 | -80.6751
1.8332 | -80.6734
1.8335 | -80.672
1.8338 | -80.6702
1.8341 | -80.6683
1.8345 | -80.6666
1.8348 | -80.6653
1.8351 | -80.6647
1.8354 | -80.6687
1.8357 | -80.6682
1.8361 | -80.6702
1.8364 | -80.6698
1.8367 | -80.6709
1.837 | -80.6721
1.8374 | -80.672
1.8377 | -80.6716
1.838 | -80.6713
1.8383 | -80.6709
1.8386 | -80.6701
1.839 | -80.6699
1.8393 | -80.6713
1.8396 | -80.6739
1.8399 | -80.6764
1.8402 | -80.6801
1.8406 | -80.6818
1.8409 | -80.6834
1.8412 | -80.6839
1.8415 | -80.6864
1.8419 | -80.6861
1.8422 | -80.6843
1.8425 | -80.6836
1.8428 | -80.6851
1.8431 | -80.687
1.8435 | -80.6892
1.8438 | -80.6893
1.8441 | -80.6917
1.8444 | -80.6941
1.8447 | -80.7003
1.8451 | -80.7028
1.8454 | -80.7052
1.8457 | -80.709
1.846 | -80.7116
1.8464 | -80.714
1.8467 | -80.7164
1.847 | -80.7188
1.8473 | -80.7207
1.8476 | -80.7227
1.848 | -80.7261
1.8483 | -80.7282
1.8486 | -80.7283
1.8489 | -80.7303
1.8492 | -80.7325
1.8496 | -80.7363
1.8499 | -80.7414
1.8502 | -80.7459
1.8505 | -80.7498
1.8509 | -80.7543
1.8512 | -80.7602
1.8515 | -80.7629
1.8518 | -80.7653
1.8521 | -80.7694
1.8525 | -80.7742
1.8528 | -80.7786
1.8531 | -80.7836
1.8534 | -80.7876
1.8537 | -80.7879
1.8541 | -80.7946
1.8544 | -80.7987
1.8547 | -80.8031
1.855 | -80.8075
1.8554 | -80.8125
1.8557 | -80.8176
1.856 | -80.8214
1.8563 | -80.8249
1.8566 | -80.8293
1.857 | -80.8339
1.8573 | -80.8389
1.8576 | -80.8447
1.8579 | -80.85
1.8582 | -80.8574
1.8586 | -80.8646
1.8589 | -80.8717
1.8592 | -80.8773
1.8595 | -80.8836
1.8598 | -80.8898
1.8602 | -80.894
1.8605 | -80.8988
1.8608 | -80.9047
1.8611 | -80.9109
1.8615 | -80.9179
1.8618 | -80.9228
1.8621 | -80.9264
1.8624 | -80.9317
1.8627 | -80.9391
1.8631 | -80.9461
1.8634 | -80.9545
1.8637 | -80.961
1.864 | -80.968
1.8643 | -80.9728
1.8647 | -80.9787
1.865 | -80.9874
1.8653 | -80.9947
1.8656 | -81.002
1.866 | -81.0103
1.8663 | -81.018
1.8666 | -81.025
1.8669 | -81.0319
1.8672 | -81.0436
1.8676 | -81.0541
1.8679 | -81.0614
1.8682 | -81.0665
1.8685 | -81.073
1.8688 | -81.0801
1.8692 | -81.0873
1.8695 | -81.0964
1.8698 | -81.1088
1.8701 | -81.1208
1.8705 | -81.1288
1.8708 | -81.1366
1.8711 | -81.1428
1.8714 | -81.1473
1.8717 | -81.1543
1.8721 | -81.1614
1.8724 | -81.1693
1.8727 | -81.1767
1.873 | -81.1858
1.8733 | -81.1943
1.8737 | -81.2033
1.874 | -81.2115
1.8743 | -81.2199
1.8746 | -81.2274
1.875 | -81.235
1.8753 | -81.2426
1.8756 | -81.2493
1.8759 | -81.2564
1.8762 | -81.2634
1.8766 | -81.2707
1.8769 | -81.2784
1.8772 | -81.2856
1.8775 | -81.2929
1.8778 | -81.3007
1.8782 | -81.3096
1.8785 | -81.3178
1.8788 | -81.3262
1.8791 | -81.3357
1.8795 | -81.3445
1.8798 | -81.355
1.8801 | -81.3646
1.8804 | -81.3721
1.8807 | -81.3804
1.8811 | -81.3873
1.8814 | -81.3931
1.8817 | -81.3976
1.882 | -81.3977
1.8823 | -81.4058
1.8827 | -81.4134
1.883 | -81.419
1.8833 | -81.4281
1.8836 | -81.4395
1.8839 | -81.4458
1.8843 | -81.4529
1.8846 | -81.461
1.8849 | -81.4702
1.8852 | -81.4783
1.8856 | -81.4853
1.8859 | -81.4921
1.8862 | -81.4997
1.8865 | -81.5081
1.8868 | -81.5136
1.8872 | -81.5203
1.8875 | -81.5275
1.8878 | -81.5351
1.8881 | -81.5423
1.8884 | -81.5506
1.8888 | -81.5593
1.8891 | -81.5682
1.8894 | -81.5749
1.8897 | -81.5825
1.8901 | -81.5911
1.8904 | -81.6002
1.8907 | -81.6078
1.891 | -81.6137
1.8913 | -81.6199
1.8917 | -81.6266
1.892 | -81.6343
1.8923 | -81.6437
1.8926 | -81.6552
1.8929 | -81.6653
1.8933 | -81.6716
1.8936 | -81.6782
1.8939 | -81.6834
1.8942 | -81.6883
1.8946 | -81.692
1.8949 | -81.6981
1.8952 | -81.7034
1.8955 | -81.7099
1.8958 | -81.7201
1.8962 | -81.7301
1.8965 | -81.7412
1.8968 | -81.7501
1.8971 | -81.7581
1.8974 | -81.7653
1.8978 | -81.772
1.8981 | -81.7768
1.8984 | -81.7786
1.8987 | -81.7829
1.8991 | -81.7866
1.8994 | -81.7889
1.8997 | -81.7938
1.9 | -81.7975
1.9003 | -81.7986
1.9007 | -81.8012
1.901 | -81.8046
1.9013 | -81.8099
1.9016 | -81.812
1.9019 | -81.8197
1.9023 | -81.827
1.9026 | -81.8315
1.9029 | -81.8379
1.9032 | -81.8422
1.9036 | -81.8451
1.9039 | -81.8472
1.9042 | -81.8504
1.9045 | -81.854
1.9048 | -81.8569
1.9052 | -81.8599
1.9055 | -81.8625
1.9058 | -81.8662
1.9061 | -81.8699
1.9064 | -81.8759
1.9068 | -81.8848
1.9071 | -81.892
1.9074 | -81.8967
1.9077 | -81.9002
1.908 | -81.9028
1.9084 | -81.9051
1.9087 | -81.9077
1.909 | -81.9119
1.9093 | -81.913
1.9097 | -81.9138
1.91 | -81.9161
1.9103 | -81.9178
1.9106 | -81.9198
1.9109 | -81.9203
1.9113 | -81.9223
1.9116 | -81.9243
1.9119 | -81.9267
1.9122 | -81.9298
1.9125 | -81.932
1.9129 | -81.9332
1.9132 | -81.9358
1.9135 | -81.9397
1.9138 | -81.9418
1.9142 | -81.9456
1.9145 | -81.9486
1.9148 | -81.9502
1.9151 | -81.9522
1.9154 | -81.9541
1.9158 | -81.9569
1.9161 | -81.9592
1.9164 | -81.9565
1.9167 | -81.9529
1.917 | -81.9535
1.9174 | -81.9575
1.9177 | -81.9617
1.918 | -81.9629
1.9183 | -81.9623
1.9187 | -81.963
1.919 | -81.962
1.9193 | -81.9632
1.9196 | -81.9644
1.9199 | -81.9656
1.9203 | -81.9661
1.9206 | -81.967
1.9209 | -81.9708
1.9212 | -81.9706
1.9215 | -81.9654
1.9219 | -81.9659
1.9222 | -81.9684
1.9225 | -81.9692
1.9228 | -81.9698
1.9232 | -81.9699
1.9235 | -81.9706
1.9238 | -81.9681
1.9241 | -81.9653
1.9244 | -81.9624
1.9248 | -81.9585
1.9251 | -81.958
1.9254 | -81.9608
1.9257 | -81.9614
1.926 | -81.9604
1.9264 | -81.957
1.9267 | -81.9558
1.927 | -81.9565
1.9273 | -81.9562
1.9277 | -81.9554
1.928 | -81.9555
1.9283 | -81.9537
1.9286 | -81.9511
1.9289 | -81.9489
1.9293 | -81.9478
1.9296 | -81.9456
1.9299 | -81.9441
1.9302 | -81.9436
1.9305 | -81.9473
1.9309 | -81.9515
1.9312 | -81.9537
1.9315 | -81.9514
1.9318 | -81.9444
1.9321 | -81.9372
1.9325 | -81.9309
1.9328 | -81.9278
1.9331 | -81.9248
1.9334 | -81.9222
1.9338 | -81.9177
1.9341 | -81.9157
1.9344 | -81.9179
1.9347 | -81.9174
1.935 | -81.9154
1.9354 | -81.9121
1.9357 | -81.9088
1.936 | -81.9057
1.9363 | -81.9025
1.9366 | -81.9007
1.937 | -81.901
1.9373 | -81.9032
1.9376 | -81.8987
1.9379 | -81.9029
1.9383 | -81.9014
1.9386 | -81.897
1.9389 | -81.8922
1.9392 | -81.8889
1.9395 | -81.8855
1.9399 | -81.8811
1.9402 | -81.8767
1.9405 | -81.8732
1.9408 | -81.8698
1.9411 | -81.8684
1.9415 | -81.8678
1.9418 | -81.8674
1.9421 | -81.8641
1.9424 | -81.8603
1.9428 | -81.8564
1.9431 | -81.8534
1.9434 | -81.8512
1.9437 | -81.8487
1.944 | -81.8456
1.9444 | -81.8396
1.9447 | -81.837
1.945 | -81.833
1.9453 | -81.829
1.9456 | -81.8257
1.946 | -81.8238
1.9463 | -81.8231
1.9466 | -81.8199
1.9469 | -81.818
1.9473 | -81.8163
1.9476 | -81.8127
1.9479 | -81.8071
1.9482 | -81.8036
1.9485 | -81.8
1.9489 | -81.795
1.9492 | -81.7898
1.9495 | -81.7859
1.9498 | -81.7815
1.9501 | -81.7767
1.9505 | -81.7709
1.9508 | -81.7668
1.9511 | -81.7637
1.9514 | -81.7632
1.9517 | -81.7626
1.9521 | -81.7604
1.9524 | -81.7546
1.9527 | -81.7516
1.953 | -81.7482
1.9534 | -81.7448
1.9537 | -81.7424
1.954 | -81.7401
1.9543 | -81.7375
1.9546 | -81.7336
1.955 | -81.7297
1.9553 | -81.7247
1.9556 | -81.7211
1.9559 | -81.7197
1.9562 | -81.7197
1.9566 | -81.7219
1.9569 | -81.7178
1.9572 | -81.7139
1.9575 | -81.7109
1.9579 | -81.7064
1.9582 | -81.7034
1.9585 | -81.6992
1.9588 | -81.6942
1.9591 | -81.6895
1.9595 | -81.6852
1.9598 | -81.6817
1.9601 | -81.6771
1.9604 | -81.674
1.9607 | -81.671
1.9611 | -81.6687
1.9614 | -81.6651
1.9617 | -81.6605
1.962 | -81.6583
1.9624 | -81.6559
1.9627 | -81.6534
1.963 | -81.6507
1.9633 | -81.6499
1.9636 | -81.6462
1.964 | -81.6428
1.9643 | -81.64
1.9646 | -81.6384
1.9649 | -81.6361
1.9652 | -81.6353
1.9656 | -81.6325
1.9659 | -81.6299
1.9662 | -81.6265
1.9665 | -81.6251
1.9669 | -81.6246
1.9672 | -81.6221
1.9675 | -81.6182
1.9678 | -81.612
1.9681 | -81.6084
1.9685 | -81.6054
1.9688 | -81.6024
1.9691 | -81.6001
1.9694 | -81.5979
1.9697 | -81.5948
1.9701 | -81.5906
1.9704 | -81.5862
1.9707 | -81.5834
1.971 | -81.5806
1.9713 | -81.5789
1.9717 | -81.5767
1.972 | -81.5723
1.9723 | -81.5662
1.9726 | -81.5631
1.973 | -81.562
1.9733 | -81.5592
1.9736 | -81.5565
1.9739 | -81.5561
1.9742 | -81.5546
1.9746 | -81.5541
1.9749 | -81.5529
1.9752 | -81.5536
1.9755 | -81.5533
1.9758 | -81.5517
1.9762 | -81.5504
1.9765 | -81.551
1.9768 | -81.5518
1.9771 | -81.5504
1.9775 | -81.5481
1.9778 | -81.5432
1.9781 | -81.5385
1.9784 | -81.5358
1.9787 | -81.5345
1.9791 | -81.5347
1.9794 | -81.5344
1.9797 | -81.5333
1.98 | -81.5329
1.9803 | -81.5326
1.9807 | -81.5314
1.981 | -81.5311
1.9813 | -81.5314
1.9816 | -81.5308
1.982 | -81.5304
1.9823 | -81.5276
1.9826 | -81.5273
1.9829 | -81.5268
1.9832 | -81.5257
1.9836 | -81.5249
1.9839 | -81.5266
1.9842 | -81.5259
1.9845 | -81.5242
1.9848 | -81.5236
1.9852 | -81.5251
1.9855 | -81.5261
1.9858 | -81.5241
1.9861 | -81.5213
1.9865 | -81.521
1.9868 | -81.521
1.9871 | -81.5211
1.9874 | -81.5251
1.9877 | -81.5259
1.9881 | -81.5244
1.9884 | -81.5242
1.9887 | -81.5238
1.989 | -81.5239
1.9893 | -81.5232
1.9897 | -81.5231
1.99 | -81.5218
1.9903 | -81.5213
1.9906 | -81.523
1.991 | -81.5241
1.9913 | -81.5227
1.9916 | -81.5221
1.9919 | -81.5223
1.9922 | -81.5211
1.9926 | -81.5224
1.9929 | -81.523
1.9932 | -81.5225
1.9935 | -81.5219
1.9938 | -81.5208
1.9942 | -81.5191
1.9945 | -81.5214
1.9948 | -81.5224
1.9951 | -81.5226
1.9954 | -81.5231
1.9958 | -81.5234
1.9961 | -81.5256
1.9964 | -81.5289
1.9967 | -81.53
1.9971 | -81.5303
1.9974 | -81.5305
1.9977 | -81.5309
1.998 | -81.5273
1.9983 | -81.5221
1.9987 | -81.5219
1.999 | -81.5183
1.9993 | -81.5139
1.9996 | -81.5133
1.9999 | -81.5122
2.0003 | -81.5112
2.0006 | -81.5109
2.0009 | -81.5084
2.0012 | -81.5041
2.0016 | -81.5046
2.0019 | -81.5044
2.0022 | -81.5043
2.0025 | -81.5026
2.0028 | -81.5012
2.0032 | -81.5027
2.0035 | -81.5018
2.0038 | -81.5001
2.0041 | -81.4997
2.0044 | -81.4987
2.0048 | -81.4983
2.0051 | -81.4973
2.0054 | -81.4956
2.0057 | -81.4942
2.0061 | -81.4928
2.0064 | -81.4935
2.0067 | -81.4928
2.007 | -81.492
2.0073 | -81.4909
2.0077 | -81.4902
2.008 | -81.4915
2.0083 | -81.4923
2.0086 | -81.4933
2.0089 | -81.4931
2.0093 | -81.4907
2.0096 | -81.4835
2.0099 | -81.4789
2.0102 | -81.4755
2.0106 | -81.4738
2.0109 | -81.4731
2.0112 | -81.4718
2.0115 | -81.471
2.0118 | -81.4704
2.0122 | -81.4688
2.0125 | -81.4682
2.0128 | -81.4678
2.0131 | -81.4679
2.0134 | -81.4682
2.0138 | -81.4659
2.0141 | -81.4637
2.0144 | -81.4628
2.0147 | -81.4608
2.0151 | -81.4605
2.0154 | -81.4582
2.0157 | -81.4578
2.016 | -81.4554
2.0163 | -81.4504
2.0167 | -81.4502
2.017 | -81.4507
2.0173 | -81.4525
2.0176 | -81.4552
2.0179 | -81.4539
2.0183 | -81.4524
2.0186 | -81.4503
2.0189 | -81.4504
2.0192 | -81.4495
2.0196 | -81.4481
2.0199 | -81.4477
2.0202 | -81.4464
2.0205 | -81.4448
2.0208 | -81.4439
2.0212 | -81.443
2.0215 | -81.4428
2.0218 | -81.4426
2.0221 | -81.4412
2.0224 | -81.4387
2.0228 | -81.4363
2.0231 | -81.4352
2.0234 | -81.4332
2.0237 | -81.4296
2.024 | -81.4222
2.0244 | -81.42
2.0247 | -81.4156
2.025 | -81.4118
2.0253 | -81.408
2.0257 | -81.4054
2.026 | -81.4039
2.0263 | -81.4039
2.0266 | -81.4021
2.0269 | -81.3994
2.0273 | -81.3979
2.0276 | -81.3953
2.0279 | -81.3934
2.0282 | -81.3916
2.0285 | -81.3895
2.0289 | -81.388
2.0292 | -81.3862
2.0295 | -81.3844
2.0298 | -81.3825
2.0302 | -81.381
2.0305 | -81.379
2.0308 | -81.3768
2.0311 | -81.374
2.0314 | -81.3726
2.0318 | -81.3652
2.0321 | -81.3617
2.0324 | -81.3623
2.0327 | -81.3639
2.033 | -81.3624
2.0334 | -81.3594
2.0337 | -81.357
2.034 | -81.352
2.0343 | -81.3496
2.0347 | -81.3479
2.035 | -81.3504
2.0353 | -81.353
2.0356 | -81.3512
2.0359 | -81.3503
2.0363 | -81.351
2.0366 | -81.3481
2.0369 | -81.3463
2.0372 | -81.3451
2.0375 | -81.3409
2.0379 | -81.3383
2.0382 | -81.3368
2.0385 | -81.333
2.0388 | -81.3308
2.0392 | -81.3297
2.0395 | -81.3239
2.0398 | -81.3215
2.0401 | -81.3192
2.0404 | -81.3163
2.0408 | -81.3173
2.0411 | -81.3175
2.0414 | -81.3155
2.0417 | -81.3104
2.042 | -81.3073
2.0424 | -81.3051
2.0427 | -81.3011
2.043 | -81.2979
2.0433 | -81.2951
2.0436 | -81.2928
2.044 | -81.2908
2.0443 | -81.2872
2.0446 | -81.2835
2.0449 | -81.2808
2.0453 | -81.2783
2.0456 | -81.276
2.0459 | -81.2739
2.0462 | -81.2717
2.0465 | -81.2687
2.0469 | -81.2659
2.0472 | -81.2642
2.0475 | -81.2611
2.0478 | -81.2577
2.0481 | -81.2556
2.0485 | -81.2538
2.0488 | -81.2536
2.0491 | -81.2507
2.0494 | -81.2452
2.0498 | -81.2412
2.0501 | -81.2374
2.0504 | -81.2349
2.0507 | -81.2334
2.051 | -81.2323
2.0514 | -81.2305
2.0517 | -81.2289
2.052 | -81.2278
2.0523 | -81.2232
2.0526 | -81.2199
2.053 | -81.2173
2.0533 | -81.214
2.0536 | -81.2125
2.0539 | -81.2101
2.0543 | -81.2082
2.0546 | -81.2096
2.0549 | -81.2114
2.0552 | -81.2101
2.0555 | -81.209
2.0559 | -81.2081
2.0562 | -81.2069
2.0565 | -81.2075
2.0568 | -81.2088
2.0571 | -81.2091
2.0575 | -81.2069
2.0578 | -81.2049
2.0581 | -81.2073
2.0584 | -81.207
2.0588 | -81.2067
2.0591 | -81.2038
2.0594 | -81.201
2.0597 | -81.1997
2.06 | -81.2014
2.0604 | -81.2045
2.0607 | -81.204
2.061 | -81.2027
2.0613 | -81.2018
2.0616 | -81.2008
2.062 | -81.1995
2.0623 | -81.1982
2.0626 | -81.1961
2.0629 | -81.1962
2.0633 | -81.198
2.0636 | -81.2001
2.0639 | -81.2004
2.0642 | -81.2
2.0645 | -81.199
2.0649 | -81.1997
2.0652 | -81.2023
2.0655 | -81.2033
2.0658 | -81.203
2.0661 | -81.2029
2.0665 | -81.2033
2.0668 | -81.2035
2.0671 | -81.2028
2.0674 | -81.2068
2.0677 | -81.2077
2.0681 | -81.2106
2.0684 | -81.2115
2.0687 | -81.2104
2.069 | -81.2086
2.0694 | -81.2056
2.0697 | -81.2041
2.07 | -81.2027
2.0703 | -81.2006
2.0706 | -81.2005
2.071 | -81.2033
2.0713 | -81.208
2.0716 | -81.213
2.0719 | -81.2194
2.0722 | -81.2245
2.0726 | -81.2253
2.0729 | -81.2258
2.0732 | -81.2278
2.0735 | -81.2287
2.0739 | -81.2281
2.0742 | -81.2287
2.0745 | -81.2298
2.0748 | -81.2309
2.0751 | -81.2325
2.0755 | -81.2346
2.0758 | -81.2363
2.0761 | -81.2375
2.0764 | -81.2419
2.0767 | -81.2421
2.0771 | -81.2432
2.0774 | -81.2464
2.0777 | -81.2483
2.078 | -81.2505
2.0784 | -81.252
2.0787 | -81.2542
2.079 | -81.2553
2.0793 | -81.2521
2.0796 | -81.2514
2.08 | -81.2547
2.0803 | -81.2577
2.0806 | -81.2578
2.0809 | -81.2588
2.0812 | -81.2605
2.0816 | -81.2627
2.0819 | -81.2647
2.0822 | -81.2667
2.0825 | -81.2691
2.0829 | -81.2707
2.0832 | -81.272
2.0835 | -81.2727
2.0838 | -81.274
2.0841 | -81.2766
2.0845 | -81.2798
2.0848 | -81.2818
2.0851 | -81.285
2.0854 | -81.2888
2.0857 | -81.2907
2.0861 | -81.2922
2.0864 | -81.2941
2.0867 | -81.2963
2.087 | -81.2996
2.0873 | -81.3015
2.0877 | -81.304
2.088 | -81.3068
2.0883 | -81.3077
2.0886 | -81.3093
2.089 | -81.3107
2.0893 | -81.3131
2.0896 | -81.3156
2.0899 | -81.3173
2.0902 | -81.3193
2.0906 | -81.3232
2.0909 | -81.325
2.0912 | -81.3282
2.0915 | -81.331
2.0918 | -81.3316
2.0922 | -81.3322
2.0925 | -81.3338
2.0928 | -81.335
2.0931 | -81.3381
2.0935 | -81.341
2.0938 | -81.3444
2.0941 | -81.3472
2.0944 | -81.3467
2.0947 | -81.3445
2.0951 | -81.3452
2.0954 | -81.341
2.0957 | -81.3373
2.096 | -81.3391
2.0963 | -81.3424
2.0967 | -81.3434
2.097 | -81.3434
2.0973 | -81.3428
2.0976 | -81.3437
2.098 | -81.347
2.0983 | -81.3511
2.0986 | -81.3531
2.0989 | -81.3543
2.0992 | -81.3573
2.0996 | -81.3553
2.0999 | -81.3567
2.1002 | -81.3604
2.1005 | -81.3623
2.1008 | -81.3637
2.1012 | -81.3658
2.1015 | -81.3684
2.1018 | -81.3696
2.1021 | -81.3711
2.1025 | -81.372
2.1028 | -81.3724
2.1031 | -81.3717
2.1034 | -81.3711
2.1037 | -81.3722
2.1041 | -81.3721
2.1044 | -81.3735
2.1047 | -81.375
2.105 | -81.3762
2.1053 | -81.3777
2.1057 | -81.3776
2.106 | -81.3771
2.1063 | -81.378
2.1066 | -81.3791
2.107 | -81.3805
2.1073 | -81.3817
2.1076 | -81.382
2.1079 | -81.3832
2.1082 | -81.3847
2.1086 | -81.3861
2.1089 | -81.387
2.1092 | -81.3866
2.1095 | -81.388
2.1098 | -81.3927
2.1102 | -81.3948
2.1105 | -81.3964
2.1108 | -81.398
2.1111 | -81.3988
2.1115 | -81.3985
2.1118 | -81.4002
2.1121 | -81.4026
2.1124 | -81.405
2.1127 | -81.4059
2.1131 | -81.4063
2.1134 | -81.4081
2.1137 | -81.4076
2.114 | -81.409
2.1143 | -81.4116
2.1147 | -81.4153
2.115 | -81.4172
2.1153 | -81.4171
2.1156 | -81.4186
2.1159 | -81.4191
2.1163 | -81.4184
2.1166 | -81.419
2.1169 | -81.4191
2.1172 | -81.4194
2.1176 | -81.4203
2.1179 | -81.4209
2.1182 | -81.419
2.1185 | -81.4186
2.1188 | -81.4181
2.1192 | -81.4168
2.1195 | -81.4131
2.1198 | -81.4109
2.1201 | -81.4087
2.1204 | -81.4092
2.1208 | -81.4122
2.1211 | -81.4109
2.1214 | -81.4065
2.1217 | -81.4008
2.1221 | -81.396
2.1224 | -81.3924
2.1227 | -81.3904
2.123 | -81.3888
2.1233 | -81.3875
2.1237 | -81.3877
2.124 | -81.3865
2.1243 | -81.3846
2.1246 | -81.3854
2.1249 | -81.387
2.1253 | -81.3876
2.1256 | -81.3865
2.1259 | -81.3844
2.1262 | -81.3831
2.1266 | -81.382
2.1269 | -81.3807
2.1272 | -81.3782
2.1275 | -81.3723
2.1278 | -81.3699
2.1282 | -81.3691
2.1285 | -81.3673
2.1288 | -81.3653
2.1291 | -81.3639
2.1294 | -81.3624
2.1298 | -81.3567
2.1301 | -81.352
2.1304 | -81.3498
2.1307 | -81.3465
2.1311 | -81.3439
2.1314 | -81.3412
2.1317 | -81.3367
2.132 | -81.3333
2.1323 | -81.3315
2.1327 | -81.3259
2.133 | -81.3215
2.1333 | -81.3187
2.1336 | -81.3155
2.1339 | -81.313
2.1343 | -81.3104
2.1346 | -81.3076
2.1349 | -81.3059
2.1352 | -81.3018
2.1355 | -81.2996
2.1359 | -81.2982
2.1362 | -81.2953
2.1365 | -81.2917
2.1368 | -81.2873
2.1372 | -81.2871
2.1375 | -81.2849
2.1378 | -81.2832
2.1381 | -81.2817
2.1384 | -81.2788
2.1388 | -81.2762
2.1391 | -81.2735
2.1394 | -81.2709
2.1397 | -81.268
2.14 | -81.2653
2.1404 | -81.2623
2.1407 | -81.2597
2.141 | -81.2566
2.1413 | -81.2526
2.1417 | -81.2474
2.142 | -81.2411
2.1423 | -81.2427
2.1426 | -81.2442
2.1429 | -81.2407
2.1433 | -81.2411
2.1436 | -81.243
2.1439 | -81.2398
2.1442 | -81.2372
2.1445 | -81.2322
2.1449 | -81.2264
2.1452 | -81.2249
2.1455 | -81.2261
2.1458 | -81.224
2.1462 | -81.2187
2.1465 | -81.2139
2.1468 | -81.211
2.1471 | -81.2094
2.1474 | -81.2072
2.1478 | -81.2033
2.1481 | -81.2001
2.1484 | -81.1969
2.1487 | -81.1937
2.149 | -81.1907
2.1494 | -81.1853
2.1497 | -81.1814
2.15 | -81.1799
2.1503 | -81.1776
2.1507 | -81.1735
2.151 | -81.172
2.1513 | -81.1696
2.1516 | -81.1662
2.1519 | -81.1627
2.1523 | -81.157
2.1526 | -81.1539
2.1529 | -81.1504
2.1532 | -81.1466
2.1535 | -81.143
2.1539 | -81.1397
2.1542 | -81.1381
2.1545 | -81.141
2.1548 | -81.1394
2.1551 | -81.1346
2.1555 | -81.1299
2.1558 | -81.1249
2.1561 | -81.1197
2.1564 | -81.1169
2.1568 | -81.1125
2.1571 | -81.1084
2.1574 | -81.1059
2.1577 | -81.1022
2.158 | -81.0989
2.1584 | -81.0958
2.1587 | -81.0914
2.159 | -81.0885
2.1593 | -81.0861
2.1596 | -81.0786
2.16 | -81.0737
2.1603 | -81.0696
2.1606 | -81.062
2.1609 | -81.0599
2.1613 | -81.0565
2.1616 | -81.0548
2.1619 | -81.0536
2.1622 | -81.048
2.1625 | -81.0437
2.1629 | -81.0407
2.1632 | -81.038
2.1635 | -81.0367
2.1638 | -81.0331
2.1641 | -81.0313
2.1645 | -81.0265
2.1648 | -81.0171
2.1651 | -81.0125
2.1654 | -81.0091
2.1658 | -81.0058
2.1661 | -81.0021
2.1664 | -80.9981
2.1667 | -80.9946
2.167 | -80.9904
2.1674 | -80.987
2.1677 | -80.9844
2.168 | -80.9821
2.1683 | -80.9797
2.1686 | -80.975
2.169 | -80.9688
2.1693 | -80.9681
2.1696 | -80.9651
2.1699 | -80.9604
2.1703 | -80.9605
2.1706 | -80.9576
2.1709 | -80.9563
2.1712 | -80.9525
2.1715 | -80.9499
2.1719 | -80.9469
2.1722 | -80.9438
2.1725 | -80.9406
2.1728 | -80.9369
2.1731 | -80.9324
2.1735 | -80.9289
2.1738 | -80.9269
2.1741 | -80.9243
2.1744 | -80.9216
2.1748 | -80.9184
2.1751 | -80.914
2.1754 | -80.9061
2.1757 | -80.8986
2.176 | -80.8945
2.1764 | -80.8918
2.1767 | -80.889
2.177 | -80.885
2.1773 | -80.8813
2.1776 | -80.8798
2.178 | -80.8782
2.1783 | -80.8783
2.1786 | -80.8754
2.1789 | -80.8704
2.1793 | -80.8655
2.1796 | -80.8624
2.1799 | -80.8591
2.1802 | -80.8543
2.1805 | -80.8513
2.1809 | -80.8491
2.1812 | -80.8451
2.1815 | -80.8428
2.1818 | -80.8397
2.1821 | -80.8353
2.1825 | -80.8314
2.1828 | -80.8278
2.1831 | -80.8248
2.1834 | -80.8198
2.1837 | -80.8124
2.1841 | -80.8087
2.1844 | -80.8077
2.1847 | -80.8051
2.185 | -80.8012
2.1854 | -80.7987
2.1857 | -80.7965
2.186 | -80.7941
2.1863 | -80.7928
2.1866 | -80.7901
2.187 | -80.7861
2.1873 | -80.7831
2.1876 | -80.78
2.1879 | -80.7782
2.1882 | -80.7753
2.1886 | -80.7723
2.1889 | -80.7687
2.1892 | -80.7648
2.1895 | -80.7611
2.1899 | -80.7584
2.1902 | -80.7557
2.1905 | -80.7533
2.1908 | -80.7501
2.1911 | -80.7475
2.1915 | -80.7443
2.1918 | -80.7416
2.1921 | -80.7399
2.1924 | -80.7363
2.1927 | -80.7324
2.1931 | -80.7297
2.1934 | -80.727
2.1937 | -80.7246
2.194 | -80.723
2.1944 | -80.7221
2.1947 | -80.7197
2.195 | -80.7166
2.1953 | -80.714
2.1956 | -80.7114
2.196 | -80.7083
2.1963 | -80.705
2.1966 | -80.7031
2.1969 | -80.7014
2.1972 | -80.6998
2.1976 | -80.7003
2.1979 | -80.6991
2.1982 | -80.698
2.1985 | -80.6954
2.1989 | -80.6922
2.1992 | -80.6892
2.1995 | -80.6865
2.1998 | -80.6845
2.2001 | -80.6819
2.2005 | -80.682
2.2008 | -80.681
2.2011 | -80.6773
2.2014 | -80.6751
2.2017 | -80.6733
2.2021 | -80.6718
2.2024 | -80.6702
2.2027 | -80.6675
2.203 | -80.6648
2.2033 | -80.6617
2.2037 | -80.6586
2.204 | -80.6558
2.2043 | -80.6526
2.2046 | -80.6504
2.205 | -80.6478
2.2053 | -80.6456
2.2056 | -80.643
2.2059 | -80.6425
2.2062 | -80.6412
2.2066 | -80.6386
2.2069 | -80.6342
2.2072 | -80.6303
2.2075 | -80.628
2.2078 | -80.6268
2.2082 | -80.6244
2.2085 | -80.6233
2.2088 | -80.6199
2.2091 | -80.6143
2.2095 | -80.6116
2.2098 | -80.6087
2.2101 | -80.6059
2.2104 | -80.6042
2.2107 | -80.6026
2.2111 | -80.6001
2.2114 | -80.5979
2.2117 | -80.5962
2.212 | -80.5941
2.2123 | -80.5928
2.2127 | -80.5909
2.213 | -80.5888
2.2133 | -80.5868
2.2136 | -80.5836
2.214 | -80.5811
2.2143 | -80.5781
2.2146 | -80.5758
2.2149 | -80.5744
2.2152 | -80.5721
2.2156 | -80.5725
2.2159 | -80.5699
2.2162 | -80.5687
2.2165 | -80.5659
2.2168 | -80.5641
2.2172 | -80.5634
2.2175 | -80.5622
2.2178 | -80.5597
2.2181 | -80.5565
2.2185 | -80.5586
2.2188 | -80.5614
2.2191 | -80.5621
2.2194 | -80.5629
2.2197 | -80.5636
2.2201 | -80.5621
2.2204 | -80.5609
2.2207 | -80.5596
2.221 | -80.558
2.2213 | -80.5566
2.2217 | -80.5542
2.222 | -80.552
2.2223 | -80.5477
2.2226 | -80.5425
2.223 | -80.5391
2.2233 | -80.5386
2.2236 | -80.5376
2.2239 | -80.5365
2.2242 | -80.5357
2.2246 | -80.5343
2.2249 | -80.5332
2.2252 | -80.5276
2.2255 | -80.5268
2.2258 | -80.5247
2.2262 | -80.5236
2.2265 | -80.523
2.2268 | -80.5225
2.2271 | -80.5224
2.2274 | -80.5218
2.2278 | -80.5213
2.2281 | -80.5195
2.2284 | -80.5193
2.2287 | -80.5175
2.2291 | -80.5188
2.2294 | -80.5213
2.2297 | -80.5219
2.23 | -80.5205
2.2303 | -80.5201
2.2307 | -80.5212
2.231 | -80.5201
2.2313 | -80.5179
2.2316 | -80.5166
2.2319 | -80.5159
2.2323 | -80.5154
2.2326 | -80.5156
2.2329 | -80.5143
2.2332 | -80.5122
2.2336 | -80.5134
2.2339 | -80.5135
2.2342 | -80.5151
2.2345 | -80.5144
2.2348 | -80.515
2.2352 | -80.5149
2.2355 | -80.5159
2.2358 | -80.5166
2.2361 | -80.5159
2.2364 | -80.5177
2.2368 | -80.5188
2.2371 | -80.5197
2.2374 | -80.5198
2.2377 | -80.519
2.2381 | -80.5188
2.2384 | -80.5193
2.2387 | -80.5196
2.239 | -80.5202
2.2393 | -80.5197
2.2397 | -80.5188
2.24 | -80.5189
2.2403 | -80.5195
2.2406 | -80.5212
2.2409 | -80.5224
2.2413 | -80.5213
2.2416 | -80.5211
2.2419 | -80.5211
2.2422 | -80.5203
2.2426 | -80.5211
2.2429 | -80.5222
2.2432 | -80.5228
2.2435 | -80.5229
2.2438 | -80.5218
2.2442 | -80.5217
2.2445 | -80.5195
2.2448 | -80.5176
2.2451 | -80.5168
2.2454 | -80.5189
2.2458 | -80.5224
2.2461 | -80.5236
2.2464 | -80.5253
2.2467 | -80.5286
2.2471 | -80.531
2.2474 | -80.5316
2.2477 | -80.532
2.248 | -80.5324
2.2483 | -80.5322
2.2487 | -80.5331
2.249 | -80.533
2.2493 | -80.5322
2.2496 | -80.5311
2.2499 | -80.5312
2.2503 | -80.5317
2.2506 | -80.5317
2.2509 | -80.5346
2.2512 | -80.5361
2.2515 | -80.5366
2.2519 | -80.5371
2.2522 | -80.5369
2.2525 | -80.5372
2.2528 | -80.5389
2.2532 | -80.5396
2.2535 | -80.5403
2.2538 | -80.5409
2.2541 | -80.5418
2.2544 | -80.5418
2.2548 | -80.542
2.2551 | -80.5427
2.2554 | -80.5433
2.2557 | -80.5446
2.256 | -80.5449
2.2564 | -80.5452
2.2567 | -80.5462
2.257 | -80.5465
2.2573 | -80.5471
2.2577 | -80.5489
2.258 | -80.5497
2.2583 | -80.5498
2.2586 | -80.55
2.2589 | -80.5487
2.2593 | -80.5491
2.2596 | -80.5504
2.2599 | -80.5511
2.2602 | -80.5514
2.2605 | -80.552
2.2609 | -80.5522
2.2612 | -80.5527
2.2615 | -80.5538
2.2618 | -80.554
2.2622 | -80.5547
2.2625 | -80.5544
2.2628 | -80.5549
2.2631 | -80.5552
2.2634 | -80.5549
2.2638 | -80.5563
2.2641 | -80.5563
2.2644 | -80.5561
2.2647 | -80.5572
2.265 | -80.5573
2.2654 | -80.557
2.2657 | -80.5561
2.266 | -80.5558
2.2663 | -80.5552
2.2667 | -80.555
2.267 | -80.5549
2.2673 | -80.5556
2.2676 | -80.5558
2.2679 | -80.5564
2.2683 | -80.5568
2.2686 | -80.5573
2.2689 | -80.558
2.2692 | -80.5587
2.2695 | -80.5592
2.2699 | -80.5597
2.2702 | -80.5607
2.2705 | -80.5611
2.2708 | -80.5637
2.2712 | -80.5608
2.2715 | -80.5582
2.2718 | -80.559
2.2721 | -80.5598
2.2724 | -80.5623
2.2728 | -80.564
2.2731 | -80.5658
2.2734 | -80.5656
2.2737 | -80.5654
2.274 | -80.5667
2.2744 | -80.5696
2.2747 | -80.5723
2.275 | -80.5751
2.2753 | -80.5767
2.2756 | -80.5766
2.276 | -80.5759
2.2763 | -80.5765
2.2766 | -80.5762
2.2769 | -80.5757
2.2773 | -80.5744
2.2776 | -80.5749
2.2779 | -80.5739
2.2782 | -80.5731
2.2785 | -80.5737
2.2789 | -80.5734
2.2792 | -80.5737
2.2795 | -80.5738
2.2798 | -80.5737
2.2801 | -80.5726
2.2805 | -80.5733
2.2808 | -80.5735
2.2811 | -80.5748
2.2814 | -80.5755
2.2818 | -80.5752
2.2821 | -80.5758
2.2824 | -80.5761
2.2827 | -80.5766
2.283 | -80.5781
2.2834 | -80.5802
2.2837 | -80.5817
2.284 | -80.5819
2.2843 | -80.583
2.2846 | -80.5832
2.285 | -80.5825
2.2853 | -80.5843
2.2856 | -80.5853
2.2859 | -80.5848
2.2863 | -80.5838
2.2866 | -80.5849
2.2869 | -80.5844
2.2872 | -80.5831
2.2875 | -80.5847
2.2879 | -80.5861
2.2882 | -80.5844
2.2885 | -80.5808
2.2888 | -80.5809
2.2891 | -80.5813
2.2895 | -80.5807
2.2898 | -80.5821
2.2901 | -80.5821
2.2904 | -80.581
2.2908 | -80.5815
2.2911 | -80.5821
2.2914 | -80.5817
2.2917 | -80.5821
2.292 | -80.5802
2.2924 | -80.5791
2.2927 | -80.5787
2.293 | -80.5784
2.2933 | -80.5782
2.2936 | -80.5776
2.294 | -80.576
2.2943 | -80.5743
2.2946 | -80.5741
2.2949 | -80.5745
2.2952 | -80.5743
2.2956 | -80.5758
2.2959 | -80.5764
2.2962 | -80.5772
2.2965 | -80.5795
2.2969 | -80.5785
2.2972 | -80.5785
2.2975 | -80.5764
2.2978 | -80.5769
2.2981 | -80.5782
2.2985 | -80.5787
2.2988 | -80.5781
2.2991 | -80.5799
2.2994 | -80.5816
2.2997 | -80.585
2.3001 | -80.5858
2.3004 | -80.586
2.3007 | -80.5861
2.301 | -80.5881
2.3014 | -80.589
2.3017 | -80.5884
2.302 | -80.5878
2.3023 | -80.5884
2.3026 | -80.5892
2.303 | -80.5903
2.3033 | -80.5919
2.3036 | -80.5908
2.3039 | -80.5884
2.3042 | -80.5866
2.3046 | -80.5859
2.3049 | -80.5865
2.3052 | -80.5871
2.3055 | -80.589
2.3059 | -80.5898
2.3062 | -80.5901
2.3065 | -80.5912
2.3068 | -80.5918
2.3071 | -80.5908
2.3075 | -80.5941
2.3078 | -80.5935
2.3081 | -80.5944
2.3084 | -80.5938
2.3087 | -80.5901
2.3091 | -80.5896
2.3094 | -80.5892
2.3097 | -80.5897
2.31 | -80.5892
2.3104 | -80.5904
2.3107 | -80.5946
2.311 | -80.5953
2.3113 | -80.5972
2.3116 | -80.5998
2.312 | -80.5999
2.3123 | -80.6014
2.3126 | -80.6028
2.3129 | -80.6023
2.3132 | -80.6029
2.3136 | -80.604
2.3139 | -80.604
2.3142 | -80.6032
2.3145 | -80.6037
2.3149 | -80.605
2.3152 | -80.6102
2.3155 | -80.6117
2.3158 | -80.6116
2.3161 | -80.6135
2.3165 | -80.6138
2.3168 | -80.6108
2.3171 | -80.6099
2.3174 | -80.61
2.3177 | -80.6125
2.3181 | -80.6135
2.3184 | -80.616
2.3187 | -80.6171
2.319 | -80.6153
2.3193 | -80.6145
2.3197 | -80.6151
2.32 | -80.6158
2.3203 | -80.616
2.3206 | -80.6174
2.321 | -80.6235
2.3213 | -80.6253
2.3216 | -80.6244
2.3219 | -80.6223
2.3222 | -80.6232
2.3226 | -80.6247
2.3229 | -80.6261
2.3232 | -80.6272
2.3235 | -80.6278
2.3238 | -80.628
2.3242 | -80.6281
2.3245 | -80.6279
2.3248 | -80.6306
2.3251 | -80.6343
2.3255 | -80.6317
2.3258 | -80.627
2.3261 | -80.6252
2.3264 | -80.6276
2.3267 | -80.629
2.3271 | -80.6295
2.3274 | -80.6269
2.3277 | -80.6252
2.328 | -80.6275
2.3283 | -80.6253
2.3287 | -80.6203
2.329 | -80.6192
2.3293 | -80.6195
2.3296 | -80.6205
2.33 | -80.6197
2.3303 | -80.6171
2.3306 | -80.6124
2.3309 | -80.6112
2.3312 | -80.6105
2.3316 | -80.6109
2.3319 | -80.6112
2.3322 | -80.6144
2.3325 | -80.6179
2.3328 | -80.6186
2.3332 | -80.6185
2.3335 | -80.6182
2.3338 | -80.6194
2.3341 | -80.6202
2.3345 | -80.6199
2.3348 | -80.6191
2.3351 | -80.6177
2.3354 | -80.6154
2.3357 | -80.6165
2.3361 | -80.615
2.3364 | -80.617
2.3367 | -80.6178
2.337 | -80.6148
2.3373 | -80.6147
2.3377 | -80.6145
2.338 | -80.6129
2.3383 | -80.6124
2.3386 | -80.6124
2.339 | -80.6106
2.3393 | -80.6092
2.3396 | -80.609
2.3399 | -80.6079
2.3402 | -80.6071
2.3406 | -80.6064
2.3409 | -80.6074
2.3412 | -80.6086
2.3415 | -80.6076
2.3418 | -80.6064
2.3422 | -80.6062
2.3425 | -80.6065
2.3428 | -80.6045
2.3431 | -80.6024
2.3434 | -80.6006
2.3438 | -80.5992
2.3441 | -80.5994
2.3444 | -80.5991
2.3447 | -80.597
2.3451 | -80.5962
2.3454 | -80.5937
2.3457 | -80.5921
2.346 | -80.5912
2.3463 | -80.5895
2.3467 | -80.5859
2.347 | -80.5835
2.3473 | -80.5835
2.3476 | -80.5829
2.3479 | -80.5799
2.3483 | -80.576
2.3486 | -80.573
2.3489 | -80.5688
2.3492 | -80.5651
2.3496 | -80.5632
2.3499 | -80.5617
2.3502 | -80.555
2.3505 | -80.5472
2.3508 | -80.5445
2.3512 | -80.5413
2.3515 | -80.5389
2.3518 | -80.5359
2.3521 | -80.5336
2.3524 | -80.5314
2.3528 | -80.5288
2.3531 | -80.5277
2.3534 | -80.5264
2.3537 | -80.5237
2.3541 | -80.5215
2.3544 | -80.5156
2.3547 | -80.5001
2.355 | -80.4959
2.3553 | -80.4934
2.3557 | -80.4912
2.356 | -80.4869
2.3563 | -80.4788
2.3566 | -80.4749
2.3569 | -80.4706
2.3573 | -80.4671
2.3576 | -80.4629
2.3579 | -80.4591
2.3582 | -80.4534
2.3586 | -80.4479
2.3589 | -80.4435
2.3592 | -80.4385
2.3595 | -80.4339
2.3598 | -80.4322
2.3602 | -80.4283
2.3605 | -80.4242
2.3608 | -80.4205
2.3611 | -80.4173
2.3614 | -80.413
2.3618 | -80.4101
2.3621 | -80.4073
2.3624 | -80.4039
2.3627 | -80.3995
2.363 | -80.3948
2.3634 | -80.3892
2.3637 | -80.3846
2.364 | -80.376
2.3643 | -80.371
2.3647 | -80.366
2.365 | -80.3612
2.3653 | -80.3578
2.3656 | -80.3523
2.3659 | -80.347
2.3663 | -80.3403
2.3666 | -80.334
2.3669 | -80.3277
2.3672 | -80.3229
2.3675 | -80.3194
2.3679 | -80.3139
2.3682 | -80.3105
2.3685 | -80.3081
2.3688 | -80.303
2.3692 | -80.2983
2.3695 | -80.2939
2.3698 | -80.2891
2.3701 | -80.2836
2.3704 | -80.2791
2.3708 | -80.2765
2.3711 | -80.2714
2.3714 | -80.2655
2.3717 | -80.2603
2.372 | -80.2546
2.3724 | -80.2485
2.3727 | -80.2424
2.373 | -80.2376
2.3733 | -80.229
2.3737 | -80.2205
2.374 | -80.2148
2.3743 | -80.2098
2.3746 | -80.2037
2.3749 | -80.1988
2.3753 | -80.1932
2.3756 | -80.1882
2.3759 | -80.1838
2.3762 | -80.1778
2.3765 | -80.1732
2.3769 | -80.1674
2.3772 | -80.1608
2.3775 | -80.155
2.3778 | -80.1508
2.3782 | -80.1481
2.3785 | -80.143
2.3788 | -80.1368
2.3791 | -80.1309
2.3794 | -80.1239
2.3798 | -80.1162
2.3801 | -80.1101
2.3804 | -80.1033
2.3807 | -80.0978
2.381 | -80.0903
2.3814 | -80.0818
2.3817 | -80.0741
2.382 | -80.0651
2.3823 | -80.0561
2.3827 | -80.0401
2.383 | -80.0362
2.3833 | -80.0319
2.3836 | -80.0261
2.3839 | -80.0175
2.3843 | -80.0079
2.3846 | -80.0005
2.3849 | -79.9933
2.3852 | -79.9869
2.3855 | -79.9831
2.3859 | -79.9806
2.3862 | -79.9742
2.3865 | -79.969
2.3868 | -79.9623
2.3871 | -79.9564
2.3875 | -79.951
2.3878 | -79.9446
2.3881 | -79.9368
2.3884 | -79.9291
2.3888 | -79.9237
2.3891 | -79.9146
2.3894 | -79.9074
2.3897 | -79.9003
2.39 | -79.8933
2.3904 | -79.8878
2.3907 | -79.8823
2.391 | -79.8757
2.3913 | -79.869
2.3916 | -79.8641
2.392 | -79.8587
2.3923 | -79.8545
2.3926 | -79.8494
2.3929 | -79.8437
2.3933 | -79.8379
2.3936 | -79.8313
2.3939 | -79.8267
2.3942 | -79.8198
2.3945 | -79.8132
2.3949 | -79.8074
2.3952 | -79.8008
2.3955 | -79.7939
2.3958 | -79.7878
2.3961 | -79.7797
2.3965 | -79.7708
2.3968 | -79.764
2.3971 | -79.7582
2.3974 | -79.7535
2.3978 | -79.7475
2.3981 | -79.7413
2.3984 | -79.7347
2.3987 | -79.7281
2.399 | -79.7226
2.3994 | -79.717
2.3997 | -79.7106
2.4 | -79.7037
2.4003 | -79.6989
2.4006 | -79.6951
2.401 | -79.6911
2.4013 | -79.6859
2.4016 | -79.6801
2.4019 | -79.6719
2.4023 | -79.6654
2.4026 | -79.6594
2.4029 | -79.6537
2.4032 | -79.6487
2.4035 | -79.6437
2.4039 | -79.6388
2.4042 | -79.6297
2.4045 | -79.62
2.4048 | -79.6147
2.4051 | -79.6085
2.4055 | -79.6044
2.4058 | -79.6001
2.4061 | -79.5948
2.4064 | -79.591
2.4068 | -79.585
2.4071 | -79.5797
2.4074 | -79.5743
2.4077 | -79.5689
2.408 | -79.5624
2.4084 | -79.5522
2.4087 | -79.5468
2.409 | -79.5433
2.4093 | -79.5371
2.4096 | -79.531
2.41 | -79.526
2.4103 | -79.5199
2.4106 | -79.5114
2.4109 | -79.4991
2.4112 | -79.4914
2.4116 | -79.4855
2.4119 | -79.4731
2.4122 | -79.4595
2.4125 | -79.4522
2.4129 | -79.4463
2.4132 | -79.4403
2.4135 | -79.4312
2.4138 | -79.4218
2.4141 | -79.4149
2.4145 | -79.4063
2.4148 | -79.3971
2.4151 | -79.3903
2.4154 | -79.3845
2.4157 | -79.3783
2.4161 | -79.3725
2.4164 | -79.366
2.4167 | -79.3594
2.417 | -79.3526
2.4174 | -79.3466
2.4177 | -79.3411
2.418 | -79.3357
2.4183 | -79.3288
2.4186 | -79.3227
2.419 | -79.317
2.4193 | -79.311
2.4196 | -79.3058
2.4199 | -79.2993
2.4202 | -79.2921
2.4206 | -79.2862
2.4209 | -79.2798
2.4212 | -79.27
2.4215 | -79.2582
2.4219 | -79.2518
2.4222 | -79.2466
2.4225 | -79.241
2.4228 | -79.2334
2.4231 | -79.227
2.4235 | -79.2216
2.4238 | -79.216
2.4241 | -79.2112
2.4244 | -79.2058
2.4247 | -79.2007
2.4251 | -79.1945
2.4254 | -79.1884
2.4257 | -79.1837
2.426 | -79.1783
2.4264 | -79.1729
2.4267 | -79.1682
2.427 | -79.1628
2.4273 | -79.158
2.4276 | -79.1532
2.428 | -79.1471
2.4283 | -79.1421
2.4286 | -79.1378
2.4289 | -79.1339
2.4292 | -79.1296
2.4296 | -79.1252
2.4299 | -79.121
2.4302 | -79.1152
2.4305 | -79.1088
2.4308 | -79.103
2.4312 | -79.0972
2.4315 | -79.0919
2.4318 | -79.086
2.4321 | -79.0803
2.4325 | -79.0753
2.4328 | -79.0698
2.4331 | -79.0652
2.4334 | -79.0606
2.4337 | -79.0551
2.4341 | -79.0483
2.4344 | -79.0397
2.4347 | -79.0336
2.435 | -79.0303
2.4353 | -79.0232
2.4357 | -79.014
2.436 | -79.0068
2.4363 | -79.0011
2.4366 | -78.9958
2.437 | -78.9907
2.4373 | -78.9849
2.4376 | -78.9744
2.4379 | -78.9656
2.4382 | -78.9606
2.4386 | -78.9552
2.4389 | -78.9499
2.4392 | -78.9431
2.4395 | -78.9379
2.4398 | -78.9331
2.4402 | -78.9278
2.4405 | -78.9224
2.4408 | -78.9161
2.4411 | -78.911
2.4415 | -78.9066
2.4418 | -78.9015
2.4421 | -78.8967
2.4424 | -78.8923
2.4427 | -78.8867
2.4431 | -78.8803
2.4434 | -78.8748
2.4437 | -78.8687
2.444 | -78.8624
2.4443 | -78.8569
2.4447 | -78.8514
2.445 | -78.8457
2.4453 | -78.8365
2.4456 | -78.8311
2.446 | -78.8265
2.4463 | -78.8206
2.4466 | -78.814
2.4469 | -78.8082
2.4472 | -78.8027
2.4476 | -78.7976
2.4479 | -78.7923
2.4482 | -78.7876
2.4485 | -78.7828
2.4488 | -78.7779
2.4492 | -78.7721
2.4495 | -78.7656
2.4498 | -78.7603
2.4501 | -78.7555
2.4505 | -78.7509
2.4508 | -78.7464
2.4511 | -78.7416
2.4514 | -78.7362
2.4517 | -78.7321
2.4521 | -78.728
2.4524 | -78.7238
2.4527 | -78.7186
2.453 | -78.7142
2.4533 | -78.7116
2.4537 | -78.7068
2.454 | -78.7026
2.4543 | -78.6981
2.4546 | -78.6933
2.455 | -78.6866
2.4553 | -78.6812
2.4556 | -78.6773
2.4559 | -78.6726
2.4562 | -78.6687
2.4566 | -78.6659
2.4569 | -78.6613
2.4572 | -78.6569
2.4575 | -78.6537
2.4578 | -78.6504
2.4582 | -78.6463
2.4585 | -78.6422
2.4588 | -78.6381
2.4591 | -78.6341
2.4594 | -78.6311
2.4598 | -78.6264
2.4601 | -78.6206
2.4604 | -78.614
2.4607 | -78.6026
2.4611 | -78.5946
2.4614 | -78.5901
2.4617 | -78.5858
2.462 | -78.5814
2.4623 | -78.5776
2.4627 | -78.5739
2.463 | -78.5715
2.4633 | -78.5679
2.4636 | -78.5637
2.4639 | -78.5593
2.4643 | -78.5544
2.4646 | -78.5499
2.4649 | -78.5458
2.4652 | -78.5399
2.4656 | -78.5341
2.4659 | -78.5324
2.4662 | -78.5267
2.4665 | -78.5197
2.4668 | -78.5156
2.4672 | -78.5121
2.4675 | -78.5088
2.4678 | -78.5057
2.4681 | -78.5009
2.4684 | -78.4928
2.4688 | -78.4875
2.4691 | -78.4838
2.4694 | -78.4794
2.4697 | -78.476
2.4701 | -78.4741
2.4704 | -78.4714
2.4707 | -78.4681
2.471 | -78.4627
2.4713 | -78.4561
2.4717 | -78.4518
2.472 | -78.4478
2.4723 | -78.4422
2.4726 | -78.437
2.4729 | -78.4329
2.4733 | -78.4295
2.4736 | -78.4255
2.4739 | -78.4231
2.4742 | -78.4189
2.4746 | -78.4148
2.4749 | -78.4121
2.4752 | -78.411
2.4755 | -78.4106
2.4758 | -78.408
2.4762 | -78.4052
2.4765 | -78.4024
2.4768 | -78.3996
2.4771 | -78.3964
2.4774 | -78.3932
2.4778 | -78.3912
2.4781 | -78.3882
2.4784 | -78.3845
2.4787 | -78.3812
2.479 | -78.3784
2.4794 | -78.375
2.4797 | -78.3723
2.48 | -78.3697
2.4803 | -78.3669
2.4807 | -78.3647
2.481 | -78.3635
2.4813 | -78.3613
2.4816 | -78.3585
2.4819 | -78.3541
2.4823 | -78.3496
2.4826 | -78.3475
2.4829 | -78.3451
2.4832 | -78.345
2.4835 | -78.3439
2.4839 | -78.3402
2.4842 | -78.3366
2.4845 | -78.3367
2.4848 | -78.3339
2.4852 | -78.3319
2.4855 | -78.3299
2.4858 | -78.3275
2.4861 | -78.3257
2.4864 | -78.3247
2.4868 | -78.3231
2.4871 | -78.3211
2.4874 | -78.3194
2.4877 | -78.3177
2.488 | -78.3154
2.4884 | -78.3102
2.4887 | -78.3057
2.489 | -78.303
2.4893 | -78.3002
2.4897 | -78.2935
2.49 | -78.2889
2.4903 | -78.2869
2.4906 | -78.2853
2.4909 | -78.2832
2.4913 | -78.2814
2.4916 | -78.28
2.4919 | -78.279
2.4922 | -78.2776
2.4925 | -78.2758
2.4929 | -78.2724
2.4932 | -78.272
2.4935 | -78.2721
2.4938 | -78.2736
2.4942 | -78.2726
2.4945 | -78.2713
2.4948 | -78.2702
2.4951 | -78.2691
2.4954 | -78.268
2.4958 | -78.2656
2.4961 | -78.2634
2.4964 | -78.2639
2.4967 | -78.2644
2.497 | -78.2634
2.4974 | -78.2622
2.4977 | -78.2608
2.498 | -78.2603
2.4983 | -78.2603
2.4987 | -78.2617
2.499 | -78.2614
2.4993 | -78.2602
2.4996 | -78.2558
2.4999 | -78.2539
2.5003 | -78.2532
2.5006 | -78.2514
2.5009 | -78.2505
2.5012 | -78.2497
2.5015 | -78.2476
2.5019 | -78.243
2.5022 | -78.2413
2.5025 | -78.24
2.5028 | -78.2393
2.5031 | -78.2404
2.5035 | -78.2411
2.5038 | -78.2414
2.5041 | -78.2409
2.5044 | -78.2397
2.5048 | -78.2387
2.5051 | -78.238
2.5054 | -78.2361
2.5057 | -78.2371
2.506 | -78.2351
2.5064 | -78.2353
2.5067 | -78.236
2.507 | -78.2361
2.5073 | -78.2359
2.5076 | -78.2363
2.508 | -78.2362
2.5083 | -78.2357
2.5086 | -78.2355
2.5089 | -78.2357
2.5093 | -78.2335
2.5096 | -78.2329
2.5099 | -78.2341
2.5102 | -78.2329
2.5105 | -78.2298
2.5109 | -78.2288
2.5112 | -78.2288
2.5115 | -78.2298
2.5118 | -78.231
2.5121 | -78.2314
2.5125 | -78.2326
2.5128 | -78.2326
2.5131 | -78.2323
2.5134 | -78.2314
2.5138 | -78.2293
2.5141 | -78.2287
2.5144 | -78.2297
2.5147 | -78.2297
2.515 | -78.2288
2.5154 | -78.2282
2.5157 | -78.2289
2.516 | -78.2308
2.5163 | -78.2322
2.5166 | -78.2331
2.517 | -78.2334
2.5173 | -78.2332
2.5176 | -78.2353
2.5179 | -78.2361
2.5183 | -78.2372
2.5186 | -78.2384
2.5189 | -78.2369
2.5192 | -78.2361
2.5195 | -78.238
2.5199 | -78.2417
2.5202 | -78.2425
2.5205 | -78.2434
2.5208 | -78.2432
2.5211 | -78.2439
2.5215 | -78.2448
2.5218 | -78.2437
2.5221 | -78.2432
2.5224 | -78.2443
2.5227 | -78.2449
2.5231 | -78.2456
2.5234 | -78.2458
2.5237 | -78.2454
2.524 | -78.2486
2.5244 | -78.2527
2.5247 | -78.257
2.525 | -78.2585
2.5253 | -78.2592
2.5256 | -78.2593
2.526 | -78.2606
2.5263 | -78.2638
2.5266 | -78.2645
2.5269 | -78.2661
2.5272 | -78.2667
2.5276 | -78.2666
2.5279 | -78.2675
2.5282 | -78.2692
2.5285 | -78.2711
2.5289 | -78.2728
2.5292 | -78.2746
2.5295 | -78.2765
2.5298 | -78.2775
2.5301 | -78.2785
2.5305 | -78.2794
2.5308 | -78.2816
2.5311 | -78.2808
2.5314 | -78.2811
2.5317 | -78.2824
2.5321 | -78.2829
2.5324 | -78.2827
2.5327 | -78.2801
2.533 | -78.2797
2.5334 | -78.2817
2.5337 | -78.2836
2.534 | -78.2845
2.5343 | -78.2857
2.5346 | -78.287
2.535 | -78.2887
2.5353 | -78.2907
2.5356 | -78.292
2.5359 | -78.2931
2.5362 | -78.295
2.5366 | -78.2969
2.5369 | -78.3002
2.5372 | -78.3002
2.5375 | -78.3002
2.5379 | -78.3017
2.5382 | -78.3029
2.5385 | -78.3042
2.5388 | -78.3052
2.5391 | -78.3068
2.5395 | -78.3085
2.5398 | -78.3099
2.5401 | -78.3121
2.5404 | -78.3117
2.5407 | -78.3124
2.5411 | -78.3136
2.5414 | -78.3123
2.5417 | -78.312
2.542 | -78.3127
2.5424 | -78.3134
2.5427 | -78.3133
2.543 | -78.3143
2.5433 | -78.3154
2.5436 | -78.3167
2.544 | -78.3185
2.5443 | -78.3184
2.5446 | -78.3203
2.5449 | -78.3228
2.5452 | -78.3242
2.5456 | -78.3253
2.5459 | -78.3265
2.5462 | -78.3278
2.5465 | -78.329
2.5469 | -78.3296
2.5472 | -78.329
2.5475 | -78.33
2.5478 | -78.3308
2.5481 | -78.3311
2.5485 | -78.332
2.5488 | -78.3327
2.5491 | -78.3341
2.5494 | -78.3343
2.5497 | -78.3327
2.5501 | -78.3319
2.5504 | -78.3331
2.5507 | -78.3344
2.551 | -78.336
2.5513 | -78.3371
2.5517 | -78.3364
2.552 | -78.3377
2.5523 | -78.3385
2.5526 | -78.3388
2.553 | -78.3381
2.5533 | -78.3402
2.5536 | -78.3439
2.5539 | -78.3477
2.5542 | -78.3521
2.5546 | -78.3557
2.5549 | -78.3572
2.5552 | -78.3554
2.5555 | -78.3541
2.5558 | -78.3562
2.5562 | -78.358
2.5565 | -78.3592
2.5568 | -78.3592
2.5571 | -78.3594
2.5575 | -78.3594
2.5578 | -78.3597
2.5581 | -78.3594
2.5584 | -78.3598
2.5587 | -78.3615
2.5591 | -78.3603
2.5594 | -78.3596
2.5597 | -78.36
2.56 | -78.3607
2.5603 | -78.3623
2.5607 | -78.3631
2.561 | -78.3626
2.5613 | -78.3648
2.5616 | -78.367
2.562 | -78.368
2.5623 | -78.3694
2.5626 | -78.3698
2.5629 | -78.3701
2.5632 | -78.3697
2.5636 | -78.3701
2.5639 | -78.3703
2.5642 | -78.3707
2.5645 | -78.3719
2.5648 | -78.3728
2.5652 | -78.3733
2.5655 | -78.3737
2.5658 | -78.3744
2.5661 | -78.3751
2.5665 | -78.3756
2.5668 | -78.3774
2.5671 | -78.3783
2.5674 | -78.3791
2.5677 | -78.3797
2.5681 | -78.3811
2.5684 | -78.3826
2.5687 | -78.3834
2.569 | -78.386
2.5693 | -78.3894
2.5697 | -78.3891
2.57 | -78.3878
2.5703 | -78.3884
2.5706 | -78.3892
2.5709 | -78.3899
2.5713 | -78.3909
2.5716 | -78.3907
2.5719 | -78.3907
2.5722 | -78.3918
2.5726 | -78.3927
2.5729 | -78.3943
2.5732 | -78.3956
2.5735 | -78.3963
2.5738 | -78.3968
2.5742 | -78.3981
2.5745 | -78.3998
2.5748 | -78.4003
2.5751 | -78.4033
2.5754 | -78.4048
2.5758 | -78.4056
2.5761 | -78.4073
2.5764 | -78.4095
2.5767 | -78.41
2.5771 | -78.4121
2.5774 | -78.4128
2.5777 | -78.4163
2.578 | -78.4205
2.5783 | -78.4214
2.5787 | -78.4227
2.579 | -78.4229
2.5793 | -78.4229
2.5796 | -78.4238
2.5799 | -78.4249
2.5803 | -78.4269
2.5806 | -78.4256
2.5809 | -78.4265
2.5812 | -78.4252
2.5816 | -78.4244
2.5819 | -78.4286
2.5822 | -78.4297
2.5825 | -78.4296
2.5828 | -78.4286
2.5832 | -78.4281
2.5835 | -78.4287
2.5838 | -78.4296
2.5841 | -78.4314
2.5844 | -78.4351
2.5848 | -78.4417
2.5851 | -78.4443
2.5854 | -78.4434
2.5857 | -78.4463
2.5861 | -78.452
2.5864 | -78.4533
2.5867 | -78.4548
2.587 | -78.4555
2.5873 | -78.456
2.5877 | -78.4587
2.588 | -78.4608
2.5883 | -78.4607
2.5886 | -78.4636
2.5889 | -78.4667
2.5893 | -78.4674
2.5896 | -78.4675
2.5899 | -78.468
2.5902 | -78.4673
2.5905 | -78.4685
2.5909 | -78.4677
2.5912 | -78.4623
2.5915 | -78.4619
2.5918 | -78.462
2.5922 | -78.4615
2.5925 | -78.4627
2.5928 | -78.4624
2.5931 | -78.4632
2.5934 | -78.464
2.5938 | -78.4645
2.5941 | -78.4653
2.5944 | -78.4668
2.5947 | -78.4676
2.595 | -78.4652
2.5954 | -78.4648
2.5957 | -78.4673
2.596 | -78.4716
2.5963 | -78.4763
2.5967 | -78.4769
2.597 | -78.4775
2.5973 | -78.4776
2.5976 | -78.4788
2.5979 | -78.4772
2.5983 | -78.477
2.5986 | -78.4776
2.5989 | -78.4779
2.5992 | -78.4785
2.5995 | -78.4795
2.5999 | -78.4798
2.6002 | -78.4803
2.6005 | -78.4797
2.6008 | -78.48
2.6012 | -78.481
2.6015 | -78.4815
2.6018 | -78.4823
2.6021 | -78.4824
2.6024 | -78.4834
2.6028 | -78.4841
2.6031 | -78.484
2.6034 | -78.4834
2.6037 | -78.482
2.604 | -78.4825
2.6044 | -78.4823
2.6047 | -78.4826
2.605 | -78.4851
2.6053 | -78.4847
2.6057 | -78.4844
2.606 | -78.4849
2.6063 | -78.4856
2.6066 | -78.4867
2.6069 | -78.4868
2.6073 | -78.4868
2.6076 | -78.4875
2.6079 | -78.4884
2.6082 | -78.4899
2.6085 | -78.4908
2.6089 | -78.492
2.6092 | -78.4935
2.6095 | -78.4954
2.6098 | -78.4963
2.6102 | -78.4977
2.6105 | -78.4996
2.6108 | -78.5037
2.6111 | -78.5079
2.6114 | -78.509
2.6118 | -78.5087
2.6121 | -78.5086
2.6124 | -78.5099
2.6127 | -78.5129
2.613 | -78.5182
2.6134 | -78.522
2.6137 | -78.5224
2.614 | -78.5229
2.6143 | -78.5216
2.6147 | -78.5223
2.615 | -78.5235
2.6153 | -78.5236
2.6156 | -78.5227
2.6159 | -78.5225
2.6163 | -78.523
2.6166 | -78.5232
2.6169 | -78.5242
2.6172 | -78.5224
2.6175 | -78.5214
2.6179 | -78.5216
2.6182 | -78.5224
2.6185 | -78.5233
2.6188 | -78.5243
2.6191 | -78.5259
2.6195 | -78.5249
2.6198 | -78.5241
2.6201 | -78.5239
2.6204 | -78.5249
2.6208 | -78.5264
2.6211 | -78.5269
2.6214 | -78.5316
2.6217 | -78.533
2.622 | -78.5333
2.6224 | -78.5339
2.6227 | -78.5325
2.623 | -78.533
2.6233 | -78.5334
2.6236 | -78.5329
2.624 | -78.5335
2.6243 | -78.5342
2.6246 | -78.535
2.6249 | -78.5357
2.6253 | -78.5361
2.6256 | -78.5367
2.6259 | -78.5372
2.6262 | -78.5389
2.6265 | -78.5384
2.6269 | -78.5378
2.6272 | -78.5382
2.6275 | -78.5383
2.6278 | -78.5383
2.6281 | -78.5401
2.6285 | -78.5411
2.6288 | -78.5416
2.6291 | -78.5407
2.6294 | -78.5399
2.6298 | -78.5403
2.6301 | -78.5409
2.6304 | -78.5408
2.6307 | -78.5412
2.631 | -78.5403
2.6314 | -78.5396
2.6317 | -78.541
2.632 | -78.5392
2.6323 | -78.538
2.6326 | -78.5386
2.633 | -78.5372
2.6333 | -78.5365
2.6336 | -78.5361
2.6339 | -78.5358
2.6343 | -78.5355
2.6346 | -78.5354
2.6349 | -78.5354
2.6352 | -78.5356
2.6355 | -78.536
2.6359 | -78.5363
2.6362 | -78.536
2.6365 | -78.5367
2.6368 | -78.5365
2.6371 | -78.5383
2.6375 | -78.5398
2.6378 | -78.5411
2.6381 | -78.5471
2.6384 | -78.55
2.6387 | -78.5503
2.6391 | -78.5506
2.6394 | -78.5511
2.6397 | -78.5507
2.64 | -78.5506
2.6404 | -78.5503
2.6407 | -78.5497
2.641 | -78.5492
2.6413 | -78.5489
2.6416 | -78.55
2.642 | -78.5495
2.6423 | -78.5495
2.6426 | -78.548
2.6429 | -78.5494
2.6432 | -78.5512
2.6436 | -78.5512
2.6439 | -78.5506
2.6442 | -78.5505
2.6445 | -78.5506
2.6449 | -78.5499
2.6452 | -78.5489
2.6455 | -78.5453
2.6458 | -78.5427
2.6461 | -78.541
2.6465 | -78.5437
2.6468 | -78.5447
2.6471 | -78.5443
2.6474 | -78.544
2.6477 | -78.5436
2.6481 | -78.5414
2.6484 | -78.5416
2.6487 | -78.5407
2.649 | -78.5412
2.6494 | -78.5417
2.6497 | -78.5397
2.65 | -78.5399
2.6503 | -78.5414
2.6506 | -78.5422
2.651 | -78.5411
2.6513 | -78.5381
2.6516 | -78.5368
2.6519 | -78.5362
2.6522 | -78.5341
2.6526 | -78.5334
2.6529 | -78.5338
2.6532 | -78.533
2.6535 | -78.5314
2.6539 | -78.5294
2.6542 | -78.5279
2.6545 | -78.5273
2.6548 | -78.5263
2.6551 | -78.5253
2.6555 | -78.5246
2.6558 | -78.5215
2.6561 | -78.5159
2.6564 | -78.5134
2.6567 | -78.5116
2.6571 | -78.5086
2.6574 | -78.5045
2.6577 | -78.5025
2.658 | -78.502
2.6584 | -78.5015
2.6587 | -78.5003
2.659 | -78.4993
2.6593 | -78.498
2.6596 | -78.497
2.66 | -78.4955
2.6603 | -78.4941
2.6606 | -78.4913
2.6609 | -78.4907
2.6612 | -78.4926
2.6616 | -78.492
2.6619 | -78.491
2.6622 | -78.4895
2.6625 | -78.489
2.6628 | -78.4896
2.6632 | -78.4892
2.6635 | -78.4865
2.6638 | -78.4839
2.6641 | -78.4821
2.6645 | -78.4797
2.6648 | -78.4768
2.6651 | -78.4755
2.6654 | -78.4743
2.6657 | -78.473
2.6661 | -78.4723
2.6664 | -78.4699
2.6667 | -78.4675
2.667 | -78.4659
2.6673 | -78.4649
2.6677 | -78.4644
2.668 | -78.4628
2.6683 | -78.4621
2.6686 | -78.463
2.669 | -78.4663
2.6693 | -78.4665
2.6696 | -78.4648
2.6699 | -78.4626
2.6702 | -78.4608
2.6706 | -78.4596
2.6709 | -78.4579
2.6712 | -78.4561
2.6715 | -78.4543
2.6718 | -78.4529
2.6722 | -78.4519
2.6725 | -78.4535
2.6728 | -78.4534
2.6731 | -78.4519
2.6735 | -78.449
2.6738 | -78.4446
2.6741 | -78.4417
2.6744 | -78.4394
2.6747 | -78.4364
2.6751 | -78.4353
2.6754 | -78.4343
2.6757 | -78.435
2.676 | -78.4333
2.6763 | -78.4315
2.6767 | -78.4314
2.677 | -78.4303
2.6773 | -78.4277
2.6776 | -78.4253
2.678 | -78.4234
2.6783 | -78.4204
2.6786 | -78.4181
2.6789 | -78.4163
2.6792 | -78.416
2.6796 | -78.4141
2.6799 | -78.415
2.6802 | -78.4149
2.6805 | -78.4134
2.6808 | -78.4121
2.6812 | -78.4099
2.6815 | -78.408
2.6818 | -78.4074
2.6821 | -78.4085
2.6824 | -78.406
2.6828 | -78.404
2.6831 | -78.4014
2.6834 | -78.3987
2.6837 | -78.3971
2.6841 | -78.3951
2.6844 | -78.393
2.6847 | -78.3913
2.685 | -78.3903
2.6853 | -78.3895
2.6857 | -78.3874
2.686 | -78.3867
2.6863 | -78.3851
2.6866 | -78.3853
2.6869 | -78.3835
2.6873 | -78.3807
2.6876 | -78.3789
2.6879 | -78.3775
2.6882 | -78.3754
2.6886 | -78.3724
2.6889 | -78.3714
2.6892 | -78.3674
2.6895 | -78.3639
2.6898 | -78.3613
2.6902 | -78.3591
2.6905 | -78.356
2.6908 | -78.3537
2.6911 | -78.3538
2.6914 | -78.3553
2.6918 | -78.3535
2.6921 | -78.3468
2.6924 | -78.3418
2.6927 | -78.3389
2.6931 | -78.3374
2.6934 | -78.3361
2.6937 | -78.335
2.694 | -78.3336
2.6943 | -78.3326
2.6947 | -78.3323
2.695 | -78.3331
2.6953 | -78.3316
2.6956 | -78.3297
2.6959 | -78.3282
2.6963 | -78.3262
2.6966 | -78.3246
2.6969 | -78.3232
2.6972 | -78.3216
2.6976 | -78.319
2.6979 | -78.3186
2.6982 | -78.3171
2.6985 | -78.3159
2.6988 | -78.3155
2.6992 | -78.3138
2.6995 | -78.3129
2.6998 | -78.3123
2.7001 | -78.3113
2.7004 | -78.3122
2.7008 | -78.3124
2.7011 | -78.312
2.7014 | -78.3103
2.7017 | -78.3096
2.7021 | -78.3085
2.7024 | -78.3081
2.7027 | -78.3064
2.703 | -78.3039
2.7033 | -78.303
2.7037 | -78.3025
2.704 | -78.3007
2.7043 | -78.297
2.7046 | -78.2948
2.7049 | -78.2926
2.7053 | -78.2919
2.7056 | -78.2916
2.7059 | -78.29
2.7062 | -78.2877
2.7066 | -78.286
2.7069 | -78.2854
2.7072 | -78.2828
2.7075 | -78.2819
2.7078 | -78.2815
2.7082 | -78.2806
2.7085 | -78.28
2.7088 | -78.2794
2.7091 | -78.2775
2.7094 | -78.277
2.7098 | -78.2768
2.7101 | -78.2752
2.7104 | -78.2748
2.7107 | -78.2738
2.711 | -78.2718
2.7114 | -78.2704
2.7117 | -78.2687
2.712 | -78.2674
2.7123 | -78.2666
2.7127 | -78.2644
2.713 | -78.2631
2.7133 | -78.2622
2.7136 | -78.2622
2.7139 | -78.2608
2.7143 | -78.2597
2.7146 | -78.2584
2.7149 | -78.2578
2.7152 | -78.2599
2.7155 | -78.2601
2.7159 | -78.2593
2.7162 | -78.2576
2.7165 | -78.2567
2.7168 | -78.2548
2.7172 | -78.2538
2.7175 | -78.2519
2.7178 | -78.2487
2.7181 | -78.2463
2.7184 | -78.2456
2.7188 | -78.2448
2.7191 | -78.244
2.7194 | -78.2425
2.7197 | -78.2418
2.72 | -78.2435
2.7204 | -78.2449
2.7207 | -78.2443
2.721 | -78.2438
2.7213 | -78.2433
2.7217 | -78.2429
2.722 | -78.242
2.7223 | -78.2408
2.7226 | -78.2406
2.7229 | -78.2401
2.7233 | -78.2398
2.7236 | -78.2409
2.7239 | -78.2393
2.7242 | -78.2364
2.7245 | -78.2374
2.7249 | -78.2383
2.7252 | -78.2389
2.7255 | -78.2393
2.7258 | -78.2398
2.7262 | -78.2405
2.7265 | -78.2406
2.7268 | -78.2406
2.7271 | -78.2412
2.7274 | -78.2403
2.7278 | -78.2414
2.7281 | -78.2414
2.7284 | -78.2411
2.7287 | -78.2408
2.729 | -78.2404
2.7294 | -78.2403
2.7297 | -78.2386
2.73 | -78.2374
2.7303 | -78.2371
2.7306 | -78.2352
2.731 | -78.2318
2.7313 | -78.2279
2.7316 | -78.2275
2.7319 | -78.2281
2.7323 | -78.2285
2.7326 | -78.2282
2.7329 | -78.2256
2.7332 | -78.2214
2.7335 | -78.2214
2.7339 | -78.2229
2.7342 | -78.2228
2.7345 | -78.2219
2.7348 | -78.2163
2.7351 | -78.2139
2.7355 | -78.215
2.7358 | -78.2159
2.7361 | -78.2159
2.7364 | -78.2158
2.7368 | -78.2165
2.7371 | -78.2194
2.7374 | -78.2195
2.7377 | -78.2185
2.738 | -78.2167
2.7384 | -78.2155
2.7387 | -78.2146
2.739 | -78.2138
2.7393 | -78.2145
2.7396 | -78.2154
2.74 | -78.2147
2.7403 | -78.2133
2.7406 | -78.2124
2.7409 | -78.2134
2.7413 | -78.2134
2.7416 | -78.2123
2.7419 | -78.2115
2.7422 | -78.2109
2.7425 | -78.2105
2.7429 | -78.2097
2.7432 | -78.2096
2.7435 | -78.2109
2.7438 | -78.2098
2.7441 | -78.2092
2.7445 | -78.2085
2.7448 | -78.2084
2.7451 | -78.2078
2.7454 | -78.2083
2.7458 | -78.2089
2.7461 | -78.2082
2.7464 | -78.2049
2.7467 | -78.203
2.747 | -78.2011
2.7474 | -78.199
2.7477 | -78.1984
2.748 | -78.1988
2.7483 | -78.1992
2.7486 | -78.1993
2.749 | -78.1991
2.7493 | -78.1985
2.7496 | -78.1983
2.7499 | -78.1968
2.7503 | -78.1967
2.7506 | -78.196
2.7509 | -78.1957
2.7512 | -78.195
2.7515 | -78.1961
2.7519 | -78.1949
2.7522 | -78.195
2.7525 | -78.1953
2.7528 | -78.1947
2.7531 | -78.1937
2.7535 | -78.1931
2.7538 | -78.1933
2.7541 | -78.1927
2.7544 | -78.1927
2.7547 | -78.1926
2.7551 | -78.1941
2.7554 | -78.1935
2.7557 | -78.193
2.756 | -78.1925
2.7564 | -78.1928
2.7567 | -78.1922
2.757 | -78.1925
2.7573 | -78.1932
2.7576 | -78.1932
2.758 | -78.1932
2.7583 | -78.193
2.7586 | -78.1932
2.7589 | -78.1925
2.7592 | -78.1887
2.7596 | -78.1852
2.7599 | -78.1846
2.7602 | -78.1853
2.7605 | -78.1853
2.7609 | -78.1856
2.7612 | -78.186
2.7615 | -78.1862
2.7618 | -78.186
2.7621 | -78.1855
2.7625 | -78.1841
2.7628 | -78.1833
2.7631 | -78.1827
2.7634 | -78.1803
2.7637 | -78.1803
2.7641 | -78.1804
2.7644 | -78.18
2.7647 | -78.1804
2.765 | -78.1808
2.7654 | -78.1818
2.7657 | -78.1825
2.766 | -78.1826
2.7663 | -78.1822
2.7666 | -78.1831
2.767 | -78.1808
2.7673 | -78.1824
2.7676 | -78.1839
2.7679 | -78.1821
2.7682 | -78.1815
2.7686 | -78.1817
2.7689 | -78.183
2.7692 | -78.1823
2.7695 | -78.1813
2.7699 | -78.1816
2.7702 | -78.1811
2.7705 | -78.1805
2.7708 | -78.1799
2.7711 | -78.1793
2.7715 | -78.1811
2.7718 | -78.1797
2.7721 | -78.182
2.7724 | -78.182
2.7727 | -78.1813
2.7731 | -78.1818
2.7734 | -78.1803
2.7737 | -78.1799
2.774 | -78.1776
2.7744 | -78.1774
2.7747 | -78.1764
2.775 | -78.1754
2.7753 | -78.1745
2.7756 | -78.1706
2.776 | -78.1695
2.7763 | -78.169
2.7766 | -78.1672
2.7769 | -78.1665
2.7772 | -78.165
2.7776 | -78.1659
2.7779 | -78.1704
2.7782 | -78.1682
2.7785 | -78.1651
2.7788 | -78.1647
2.7792 | -78.1631
2.7795 | -78.1632
2.7798 | -78.1616
2.7801 | -78.1573
2.7805 | -78.1513
2.7808 | -78.1484
2.7811 | -78.1475
2.7814 | -78.1465
2.7817 | -78.1463
2.7821 | -78.1475
2.7824 | -78.1472
2.7827 | -78.1456
2.783 | -78.1439
2.7833 | -78.1418
2.7837 | -78.14
2.784 | -78.1395
2.7843 | -78.1381
2.7846 | -78.1367
2.785 | -78.1346
2.7853 | -78.1359
2.7856 | -78.134
2.7859 | -78.1314
2.7862 | -78.1304
2.7866 | -78.129
2.7869 | -78.1279
2.7872 | -78.1263
2.7875 | -78.1246
2.7878 | -78.1241
2.7882 | -78.1233
2.7885 | -78.1219
2.7888 | -78.1209
2.7891 | -78.1195
2.7895 | -78.1183
2.7898 | -78.1147
2.7901 | -78.1127
2.7904 | -78.112
2.7907 | -78.1119
2.7911 | -78.1108
2.7914 | -78.1108
2.7917 | -78.1094
2.792 | -78.108
2.7923 | -78.1059
2.7927 | -78.103
2.793 | -78.1001
2.7933 | -78.0982
2.7936 | -78.0986
2.794 | -78.0985
2.7943 | -78.0977
2.7946 | -78.0966
2.7949 | -78.0959
2.7952 | -78.0955
2.7956 | -78.0944
2.7959 | -78.0921
2.7962 | -78.0896
2.7965 | -78.0867
2.7968 | -78.0837
2.7972 | -78.082
2.7975 | -78.0795
2.7978 | -78.0803
2.7981 | -78.0791
2.7984 | -78.0761
2.7988 | -78.0704
2.7991 | -78.065
2.7994 | -78.0636
2.7997 | -78.0607
2.8001 | -78.0594
2.8004 | -78.0555
2.8007 | -78.0518
2.801 | -78.0489
2.8013 | -78.0468
2.8017 | -78.0446
2.802 | -78.0435
2.8023 | -78.0414
2.8026 | -78.041
2.8029 | -78.0403
2.8033 | -78.0381
2.8036 | -78.0358
2.8039 | -78.034
2.8042 | -78.0321
2.8046 | -78.0303
2.8049 | -78.0288
2.8052 | -78.0272
2.8055 | -78.0257
2.8058 | -78.0255
2.8062 | -78.023
2.8065 | -78.0197
2.8068 | -78.0164
2.8071 | -78.0135
2.8074 | -78.0108
2.8078 | -78.008
2.8081 | -78.0066
2.8084 | -78.0059
2.8087 | -78.001
2.8091 | -77.9949
2.8094 | -77.9925
2.8097 | -77.9873
2.81 | -77.9809
2.8103 | -77.976
2.8107 | -77.9711
2.811 | -77.9676
2.8113 | -77.9652
2.8116 | -77.9646
2.8119 | -77.9618
2.8123 | -77.9599
2.8126 | -77.958
2.8129 | -77.9557
2.8132 | -77.9517
2.8136 | -77.9489
2.8139 | -77.9451
2.8142 | -77.9422
2.8145 | -77.9396
2.8148 | -77.9371
2.8152 | -77.9346
2.8155 | -77.931
2.8158 | -77.9302
2.8161 | -77.9284
2.8164 | -77.9254
2.8168 | -77.9222
2.8171 | -77.9196
2.8174 | -77.9164
2.8177 | -77.9135
2.8181 | -77.9098
2.8184 | -77.9074
2.8187 | -77.9048
2.819 | -77.8989
2.8193 | -77.8927
2.8197 | -77.8875
2.82 | -77.883
2.8203 | -77.8786
2.8206 | -77.8741
2.8209 | -77.8682
2.8213 | -77.864
2.8216 | -77.8629
2.8219 | -77.8611
2.8222 | -77.8565
2.8226 | -77.8531
2.8229 | -77.8516
2.8232 | -77.8491
2.8235 | -77.8465
2.8238 | -77.8426
2.8242 | -77.8393
2.8245 | -77.8329
2.8248 | -77.8273
2.8251 | -77.8238
2.8254 | -77.8196
2.8258 | -77.8153
2.8261 | -77.8109
2.8264 | -77.8064
2.8267 | -77.8031
2.827 | -77.8
2.8274 | -77.7931
2.8277 | -77.7895
2.828 | -77.7865
2.8283 | -77.7827
2.8287 | -77.7792
2.829 | -77.7754
2.8293 | -77.7705
2.8296 | -77.7643
2.8299 | -77.7596
2.8303 | -77.7554
2.8306 | -77.7518
2.8309 | -77.7487
2.8312 | -77.7442
2.8315 | -77.7377
2.8319 | -77.7301
2.8322 | -77.725
2.8325 | -77.7209
2.8328 | -77.7176
2.8332 | -77.7151
2.8335 | -77.7119
2.8338 | -77.7079
2.8341 | -77.704
2.8344 | -77.7006
2.8348 | -77.6969
2.8351 | -77.6945
2.8354 | -77.6868
2.8357 | -77.6791
2.836 | -77.6767
2.8364 | -77.6745
2.8367 | -77.667
2.837 | -77.6634
2.8373 | -77.66
2.8377 | -77.6557
2.838 | -77.6518
2.8383 | -77.6479
2.8386 | -77.6443
2.8389 | -77.6416
2.8393 | -77.6382
2.8396 | -77.635
2.8399 | -77.6305
2.8402 | -77.627
2.8405 | -77.624
2.8409 | -77.6186
2.8412 | -77.6135
2.8415 | -77.6026
2.8418 | -77.598
2.8422 | -77.5936
2.8425 | -77.5887
2.8428 | -77.5838
2.8431 | -77.5722
2.8434 | -77.5648
2.8438 | -77.5579
2.8441 | -77.551
2.8444 | -77.546
2.8447 | -77.5381
2.845 | -77.5332
2.8454 | -77.5238
2.8457 | -77.5161
2.846 | -77.5032
2.8463 | -77.492
2.8466 | -77.4871
2.847 | -77.4809
2.8473 | -77.4757
2.8476 | -77.4732
2.8479 | -77.4686
2.8483 | -77.4623
2.8486 | -77.4571
2.8489 | -77.4543
2.8492 | -77.444
2.8495 | -77.4261
2.8499 | -77.4178
2.8502 | -77.4112
2.8505 | -77.4014
2.8508 | -77.3941
2.8511 | -77.3882
2.8515 | -77.3829
2.8518 | -77.3762
2.8521 | -77.3694
2.8524 | -77.3649
2.8528 | -77.3598
2.8531 | -77.3509
2.8534 | -77.3365
2.8537 | -77.3288
2.854 | -77.3228
2.8544 | -77.3168
2.8547 | -77.3108
2.855 | -77.305
2.8553 | -77.297
2.8556 | -77.292
2.856 | -77.2854
2.8563 | -77.2752
2.8566 | -77.2669
2.8569 | -77.2612
2.8573 | -77.2538
2.8576 | -77.2459
2.8579 | -77.2385
2.8582 | -77.2316
2.8585 | -77.2246
2.8589 | -77.218
2.8592 | -77.2111
2.8595 | -77.2025
2.8598 | -77.1972
2.8601 | -77.1922
2.8605 | -77.1859
2.8608 | -77.1793
2.8611 | -77.1725
2.8614 | -77.166
2.8618 | -77.1593
2.8621 | -77.1517
2.8624 | -77.1432
2.8627 | -77.1321
2.863 | -77.1219
2.8634 | -77.113
2.8637 | -77.1036
2.864 | -77.095
2.8643 | -77.0877
2.8646 | -77.0768
2.865 | -77.0682
2.8653 | -77.0611
2.8656 | -77.0538
2.8659 | -77.0463
2.8662 | -77.0394
2.8666 | -77.0334
2.8669 | -77.0266
2.8672 | -77.0183
2.8675 | -77.0104
2.8679 | -77.003
2.8682 | -76.9962
2.8685 | -76.9887
2.8688 | -76.9814
2.8691 | -76.974
2.8695 | -76.9654
2.8698 | -76.9575
2.8701 | -76.9501
2.8704 | -76.9427
2.8707 | -76.935
2.8711 | -76.9282
2.8714 | -76.9211
2.8717 | -76.9125
2.872 | -76.9029
2.8724 | -76.8932
2.8727 | -76.8856
2.873 | -76.8785
2.8733 | -76.87
2.8736 | -76.8625
2.874 | -76.8541
2.8743 | -76.8468
2.8746 | -76.8395
2.8749 | -76.8318
2.8752 | -76.8237
2.8756 | -76.8165
2.8759 | -76.8103
2.8762 | -76.8035
2.8765 | -76.7967
2.8769 | -76.7889
2.8772 | -76.7812
2.8775 | -76.7737
2.8778 | -76.7664
2.8781 | -76.7582
2.8785 | -76.7504
2.8788 | -76.7426
2.8791 | -76.7355
2.8794 | -76.7284
2.8797 | -76.7199
2.8801 | -76.7118
2.8804 | -76.7044
2.8807 | -76.6973
2.881 | -76.6892
2.8814 | -76.6807
2.8817 | -76.6713
2.882 | -76.6634
2.8823 | -76.6563
2.8826 | -76.6482
2.883 | -76.6414
2.8833 | -76.6333
2.8836 | -76.6264
2.8839 | -76.619
2.8842 | -76.6131
2.8846 | -76.6062
2.8849 | -76.5983
2.8852 | -76.5904
2.8855 | -76.5803
2.8859 | -76.5719
2.8862 | -76.5665
2.8865 | -76.5603
2.8868 | -76.553
2.8871 | -76.5451
2.8875 | -76.5371
2.8878 | -76.5287
2.8881 | -76.5199
2.8884 | -76.5116
2.8887 | -76.5033
2.8891 | -76.4954
2.8894 | -76.4881
2.8897 | -76.4796
2.89 | -76.472
2.8904 | -76.4648
2.8907 | -76.4567
2.891 | -76.4488
2.8913 | -76.4411
2.8916 | -76.4344
2.892 | -76.4268
2.8923 | -76.4189
2.8926 | -76.4111
2.8929 | -76.4022
2.8932 | -76.394
2.8936 | -76.3864
2.8939 | -76.379
2.8942 | -76.3696
2.8945 | -76.3595
2.8948 | -76.3502
2.8952 | -76.3426
2.8955 | -76.3353
2.8958 | -76.3282
2.8961 | -76.3214
2.8965 | -76.3134
2.8968 | -76.3056
2.8971 | -76.2985
2.8974 | -76.2911
2.8977 | -76.2835
2.8981 | -76.2737
2.8984 | -76.263
2.8987 | -76.2544
2.899 | -76.2461
2.8993 | -76.2391
2.8997 | -76.2325
2.9 | -76.2255
2.9003 | -76.2183
2.9006 | -76.2111
2.901 | -76.2038
2.9013 | -76.1943
2.9016 | -76.1882
2.9019 | -76.1813
2.9022 | -76.1741
2.9026 | -76.1642
2.9029 | -76.1565
2.9032 | -76.1484
2.9035 | -76.1384
2.9038 | -76.1283
2.9042 | -76.1199
2.9045 | -76.1124
2.9048 | -76.1047
2.9051 | -76.0981
2.9055 | -76.0916
2.9058 | -76.0844
2.9061 | -76.0771
2.9064 | -76.0705
2.9067 | -76.0641
2.9071 | -76.0578
2.9074 | -76.0513
2.9077 | -76.0452
2.908 | -76.039
2.9083 | -76.0324
2.9087 | -76.0264
2.909 | -76.0206
2.9093 | -76.0149
2.9096 | -76.0088
2.91 | -76.0032
2.9103 | -75.9977
2.9106 | -75.9915
2.9109 | -75.984
2.9112 | -75.9747
2.9116 | -75.9683
2.9119 | -75.9633
2.9122 | -75.9574
2.9125 | -75.9511
2.9128 | -75.946
2.9132 | -75.9404
2.9135 | -75.9338
2.9138 | -75.9278
2.9141 | -75.9229
2.9144 | -75.9167
2.9148 | -75.911
2.9151 | -75.9048
2.9154 | -75.8996
2.9157 | -75.8944
2.9161 | -75.8891
2.9164 | -75.8824
2.9167 | -75.8763
2.917 | -75.8707
2.9173 | -75.8633
2.9177 | -75.8543
2.918 | -75.8484
2.9183 | -75.8434
2.9186 | -75.8366
2.9189 | -75.8297
2.9193 | -75.8244
2.9196 | -75.8194
2.9199 | -75.8149
2.9202 | -75.8129
2.9206 | -75.8095
2.9209 | -75.8054
2.9212 | -75.8006
2.9215 | -75.7956
2.9218 | -75.79
2.9222 | -75.7842
2.9225 | -75.7789
2.9228 | -75.7735
2.9231 | -75.7686
2.9234 | -75.7636
2.9238 | -75.7592
2.9241 | -75.754
2.9244 | -75.7484
2.9247 | -75.7439
2.9251 | -75.7401
2.9254 | -75.736
2.9257 | -75.7309
2.926 | -75.7254
2.9263 | -75.7215
2.9267 | -75.716
2.927 | -75.7113
2.9273 | -75.7088
2.9276 | -75.7069
2.9279 | -75.7049
2.9283 | -75.7017
2.9286 | -75.6975
2.9289 | -75.6938
2.9292 | -75.6897
2.9296 | -75.6884
2.9299 | -75.6864
2.9302 | -75.6859
2.9305 | -75.6833
2.9308 | -75.6795
2.9312 | -75.6753
2.9315 | -75.6699
2.9318 | -75.6674
2.9321 | -75.6664
2.9324 | -75.6644
2.9328 | -75.6612
2.9331 | -75.658
2.9334 | -75.6547
2.9337 | -75.6515
2.9341 | -75.6465
2.9344 | -75.6422
2.9347 | -75.6393
2.935 | -75.6371
2.9353 | -75.6353
2.9357 | -75.6322
2.936 | -75.6299
2.9363 | -75.6273
2.9366 | -75.6251
2.9369 | -75.6223
2.9373 | -75.6196
2.9376 | -75.6179
2.9379 | -75.6157
2.9382 | -75.6133
2.9385 | -75.6113
2.9389 | -75.6131
2.9392 | -75.6118
2.9395 | -75.6101
2.9398 | -75.6086
2.9402 | -75.6063
2.9405 | -75.6024
2.9408 | -75.5995
2.9411 | -75.5979
2.9414 | -75.5972
2.9418 | -75.5967
2.9421 | -75.5951
2.9424 | -75.5934
2.9427 | -75.5913
2.943 | -75.5891
2.9434 | -75.5878
2.9437 | -75.5867
2.944 | -75.5858
2.9443 | -75.5861
2.9447 | -75.5859
2.945 | -75.5857
2.9453 | -75.5856
2.9456 | -75.585
2.9459 | -75.5841
2.9463 | -75.5828
2.9466 | -75.58
2.9469 | -75.5787
2.9472 | -75.5783
2.9475 | -75.5777
2.9479 | -75.5779
2.9482 | -75.5814
2.9485 | -75.5811
2.9488 | -75.5812
2.9492 | -75.5811
2.9495 | -75.5813
2.9498 | -75.5803
2.9501 | -75.5804
2.9504 | -75.5799
2.9508 | -75.5801
2.9511 | -75.5805
2.9514 | -75.5807
2.9517 | -75.581
2.952 | -75.582
2.9524 | -75.5819
2.9527 | -75.582
2.953 | -75.5818
2.9533 | -75.5822
2.9537 | -75.5827
2.954 | -75.5839
2.9543 | -75.5839
2.9546 | -75.5848
2.9549 | -75.585
2.9553 | -75.5851
2.9556 | -75.586
2.9559 | -75.5861
2.9562 | -75.5866
2.9565 | -75.5877
2.9569 | -75.5887
2.9572 | -75.5932
2.9575 | -75.5934
2.9578 | -75.5935
2.9581 | -75.5965
2.9585 | -75.5994
2.9588 | -75.6007
2.9591 | -75.601
2.9594 | -75.6019
2.9598 | -75.6031
2.9601 | -75.6046
2.9604 | -75.6061
2.9607 | -75.6077
2.961 | -75.6094
2.9614 | -75.612
2.9617 | -75.6145
2.962 | -75.6163
2.9623 | -75.6181
2.9626 | -75.6202
2.963 | -75.622
2.9633 | -75.6241
2.9636 | -75.6263
2.9639 | -75.6279
2.9643 | -75.6297
2.9646 | -75.6313
2.9649 | -75.6333
2.9652 | -75.6352
2.9655 | -75.6374
2.9659 | -75.6402
2.9662 | -75.6434
2.9665 | -75.6488
2.9668 | -75.6514
2.9671 | -75.6538
2.9675 | -75.6569
2.9678 | -75.6594
2.9681 | -75.661
2.9684 | -75.6634
2.9688 | -75.667
2.9691 | -75.6697
2.9694 | -75.6723
2.9697 | -75.6753
2.97 | -75.6784
2.9704 | -75.6813
2.9707 | -75.6839
2.971 | -75.6864
2.9713 | -75.6894
2.9716 | -75.6926
2.972 | -75.6954
2.9723 | -75.6985
2.9726 | -75.7021
2.9729 | -75.7053
2.9733 | -75.7089
2.9736 | -75.7112
2.9739 | -75.715
2.9742 | -75.7185
2.9745 | -75.7226
2.9749 | -75.7272
2.9752 | -75.7311
2.9755 | -75.7361
2.9758 | -75.7418
2.9761 | -75.7459
2.9765 | -75.7498
2.9768 | -75.7542
2.9771 | -75.7584
2.9774 | -75.7625
2.9778 | -75.7661
2.9781 | -75.7709
2.9784 | -75.7764
2.9787 | -75.7815
2.979 | -75.7853
2.9794 | -75.7892
2.9797 | -75.7925
2.98 | -75.7967
2.9803 | -75.8026
2.9806 | -75.8091
2.981 | -75.8142
2.9813 | -75.818
2.9816 | -75.8212
2.9819 | -75.8253
2.9823 | -75.8292
2.9826 | -75.8327
2.9829 | -75.8368
2.9832 | -75.8408
2.9835 | -75.8453
2.9839 | -75.8498
2.9842 | -75.8548
2.9845 | -75.8595
2.9848 | -75.8642
2.9851 | -75.8709
2.9855 | -75.8751
2.9858 | -75.8799
2.9861 | -75.8844
2.9864 | -75.8898
2.9867 | -75.8951
2.9871 | -75.8995
2.9874 | -75.9043
2.9877 | -75.9091
2.988 | -75.9145
2.9884 | -75.9198
2.9887 | -75.9249
2.989 | -75.9298
2.9893 | -75.9348
2.9896 | -75.9395
2.99 | -75.9435
2.9903 | -75.9477
2.9906 | -75.9526
2.9909 | -75.9575
2.9912 | -75.9623
2.9916 | -75.9671
2.9919 | -75.9723
2.9922 | -75.9783
2.9925 | -75.9839
2.9929 | -75.9893
2.9932 | -75.9947
2.9935 | -76.0002
2.9938 | -76.0061
2.9941 | -76.0119
2.9945 | -76.0174
2.9948 | -76.0224
2.9951 | -76.0259
2.9954 | -76.0308
2.9957 | -76.0363
2.9961 | -76.0418
2.9964 | -76.0472
2.9967 | -76.0529
2.997 | -76.0583
2.9974 | -76.0627
2.9977 | -76.0654
2.998 | -76.0701
2.9983 | -76.0766
2.9986 | -76.0824
2.999 | -76.0889
2.9993 | -76.0947
2.9996 | -76.1002
2.9999 | -76.1062
3.0002 | -76.1141
3.0006 | -76.122
3.0009 | -76.1285
3.0012 | -76.1349
3.0015 | -76.1403
3.0019 | -76.1456
3.0022 | -76.1527
3.0025 | -76.1594
3.0028 | -76.1657
3.0031 | -76.1717
3.0035 | -76.1776
3.0038 | -76.1844
3.0041 | -76.1896
3.0044 | -76.1955
3.0047 | -76.2013
3.0051 | -76.2076
3.0054 | -76.2141
3.0057 | -76.2205
3.006 | -76.2274
3.0063 | -76.2344
3.0067 | -76.2405
3.007 | -76.246
3.0073 | -76.2524
3.0076 | -76.2585
3.008 | -76.265
3.0083 | -76.2723
3.0086 | -76.2792
3.0089 | -76.2866
3.0092 | -76.2932
3.0096 | -76.2992
3.0099 | -76.3051
3.0102 | -76.31
3.0105 | -76.3157
3.0108 | -76.3244
3.0112 | -76.3333
3.0115 | -76.3396
3.0118 | -76.3452
3.0121 | -76.3534
3.0125 | -76.3628
3.0128 | -76.3709
3.0131 | -76.3754
3.0134 | -76.3826
3.0137 | -76.3898
3.0141 | -76.3962
3.0144 | -76.4022
3.0147 | -76.4089
3.015 | -76.4151
3.0153 | -76.4218
3.0157 | -76.4279
3.016 | -76.4348
3.0163 | -76.4417
3.0166 | -76.4488
3.017 | -76.4555
3.0173 | -76.4616
3.0176 | -76.467
3.0179 | -76.4738
3.0182 | -76.4816
3.0186 | -76.4884
3.0189 | -76.4949
3.0192 | -76.5017
3.0195 | -76.509
3.0198 | -76.5162
3.0202 | -76.5234
3.0205 | -76.5301
3.0208 | -76.5374
3.0211 | -76.5443
3.0215 | -76.5505
3.0218 | -76.5567
3.0221 | -76.5615
3.0224 | -76.5713
3.0227 | -76.5798
3.0231 | -76.587
3.0234 | -76.5941
3.0237 | -76.601
3.024 | -76.6096
3.0243 | -76.6191
3.0247 | -76.6273
3.025 | -76.6342
3.0253 | -76.6406
3.0256 | -76.6477
3.0259 | -76.6547
3.0263 | -76.6613
3.0266 | -76.6677
3.0269 | -76.6745
3.0272 | -76.6823
3.0276 | -76.6897
3.0279 | -76.6988
3.0282 | -76.7073
3.0285 | -76.7142
3.0288 | -76.722
3.0292 | -76.7292
3.0295 | -76.737
3.0298 | -76.7443
3.0301 | -76.7513
3.0304 | -76.7576
3.0308 | -76.7653
3.0311 | -76.7724
3.0314 | -76.7792
3.0317 | -76.7865
3.0321 | -76.7938
3.0324 | -76.8017
3.0327 | -76.8083
3.033 | -76.8148
3.0333 | -76.8231
3.0337 | -76.8321
3.034 | -76.8401
3.0343 | -76.8475
3.0346 | -76.854
3.0349 | -76.8608
3.0353 | -76.8689
3.0356 | -76.8787
3.0359 | -76.8884
3.0362 | -76.8948
3.0366 | -76.8997
3.0369 | -76.9058
3.0372 | -76.9128
3.0375 | -76.9197
3.0378 | -76.9268
3.0382 | -76.9338
3.0385 | -76.9406
3.0388 | -76.9483
3.0391 | -76.9557
3.0394 | -76.9658
3.0398 | -76.9756
3.0401 | -76.9833
3.0404 | -76.99
3.0407 | -76.9967
3.0411 | -77.0033
3.0414 | -77.0129
3.0417 | -77.0212
3.042 | -77.0282
3.0423 | -77.0351
3.0427 | -77.0418
3.043 | -77.0488
3.0433 | -77.055
3.0436 | -77.0613
3.0439 | -77.0679
3.0443 | -77.0748
3.0446 | -77.0822
3.0449 | -77.0898
3.0452 | -77.0968
3.0456 | -77.1047
3.0459 | -77.1125
3.0462 | -77.125
3.0465 | -77.1326
3.0468 | -77.1407
3.0472 | -77.1488
3.0475 | -77.1571
3.0478 | -77.1691
3.0481 | -77.1785
3.0484 | -77.1874
3.0488 | -77.196
3.0491 | -77.2031
3.0494 | -77.2128
3.0497 | -77.2201
3.0501 | -77.2262
3.0504 | -77.2371
3.0507 | -77.246
3.051 | -77.2529
3.0513 | -77.2662
3.0517 | -77.2784
3.052 | -77.2843
3.0523 | -77.2911
3.0526 | -77.298
3.0529 | -77.3048
3.0533 | -77.3106
3.0536 | -77.318
3.0539 | -77.3258
3.0542 | -77.3313
3.0545 | -77.3423
3.0549 | -77.3596
3.0552 | -77.3677
3.0555 | -77.3736
3.0558 | -77.3828
3.0562 | -77.3904
3.0565 | -77.3963
3.0568 | -77.4024
3.0571 | -77.4095
3.0574 | -77.4166
3.0578 | -77.4225
3.0581 | -77.428
3.0584 | -77.4362
3.0587 | -77.4497
3.059 | -77.4567
3.0594 | -77.463
3.0597 | -77.4689
3.06 | -77.4748
3.0603 | -77.4799
3.0607 | -77.4849
3.061 | -77.4903
3.0613 | -77.4964
3.0616 | -77.5023
3.0619 | -77.5094
3.0623 | -77.5196
3.0626 | -77.528
3.0629 | -77.5343
3.0632 | -77.5413
3.0635 | -77.5484
3.0639 | -77.5552
3.0642 | -77.5601
3.0645 | -77.5647
3.0648 | -77.5696
3.0652 | -77.5754
3.0655 | -77.5824
3.0658 | -77.5882
3.0661 | -77.5931
3.0664 | -77.5978
3.0668 | -77.6032
3.0671 | -77.6088
3.0674 | -77.6147
3.0677 | -77.6201
3.068 | -77.6263
3.0684 | -77.6329
3.0687 | -77.6357
3.069 | -77.6452
3.0693 | -77.6539
3.0697 | -77.6606
3.07 | -77.6669
3.0703 | -77.6723
3.0706 | -77.677
3.0709 | -77.6818
3.0713 | -77.6886
3.0716 | -77.695
3.0719 | -77.7003
3.0722 | -77.7057
3.0725 | -77.7111
3.0729 | -77.7164
3.0732 | -77.722
3.0735 | -77.7276
3.0738 | -77.7337
3.0741 | -77.7401
3.0745 | -77.7454
3.0748 | -77.7505
3.0751 | -77.7548
3.0754 | -77.7594
3.0758 | -77.7641
3.0761 | -77.7705
3.0764 | -77.7757
3.0767 | -77.781
3.077 | -77.7867
3.0774 | -77.7931
3.0777 | -77.7989
3.078 | -77.8041
3.0783 | -77.8092
3.0786 | -77.8157
3.079 | -77.8227
3.0793 | -77.8299
3.0796 | -77.8352
3.0799 | -77.8402
3.0803 | -77.8456
3.0806 | -77.8492
3.0809 | -77.8526
3.0812 | -77.8567
3.0815 | -77.8616
3.0819 | -77.8657
3.0822 | -77.8702
3.0825 | -77.8761
3.0828 | -77.8811
3.0831 | -77.886
3.0835 | -77.891
3.0838 | -77.8955
3.0841 | -77.9013
3.0844 | -77.9067
3.0848 | -77.9117
3.0851 | -77.9165
3.0854 | -77.922
3.0857 | -77.9274
3.086 | -77.9329
3.0864 | -77.9382
3.0867 | -77.9421
3.087 | -77.9433
3.0873 | -77.9488
3.0876 | -77.9544
3.088 | -77.9592
3.0883 | -77.9643
3.0886 | -77.9694
3.0889 | -77.9747
3.0893 | -77.9814
3.0896 | -77.9873
3.0899 | -77.992
3.0902 | -77.9968
3.0905 | -78.0031
3.0909 | -78.0087
3.0912 | -78.0129
3.0915 | -78.0189
3.0918 | -78.0241
3.0921 | -78.0288
3.0925 | -78.0336
3.0928 | -78.0385
3.0931 | -78.0432
3.0934 | -78.0509
3.0938 | -78.0565
3.0941 | -78.0615
3.0944 | -78.0665
3.0947 | -78.0708
3.095 | -78.0757
3.0954 | -78.0807
3.0957 | -78.0863
3.096 | -78.0926
3.0963 | -78.0972
3.0966 | -78.1023
3.097 | -78.1066
3.0973 | -78.1097
3.0976 | -78.1126
3.0979 | -78.1167
3.0982 | -78.1199
3.0986 | -78.1215
3.0989 | -78.1248
3.0992 | -78.1285
3.0995 | -78.1322
3.0999 | -78.1356
3.1002 | -78.1398
3.1005 | -78.1446
3.1008 | -78.1492
3.1011 | -78.1538
3.1015 | -78.1598
3.1018 | -78.165
3.1021 | -78.1688
3.1024 | -78.1727
3.1027 | -78.1789
3.1031 | -78.186
3.1034 | -78.1925
3.1037 | -78.1969
3.104 | -78.2007
3.1044 | -78.2047
3.1047 | -78.2089
3.105 | -78.2139
3.1053 | -78.2188
3.1056 | -78.2227
3.106 | -78.2256
3.1063 | -78.2304
3.1066 | -78.236
3.1069 | -78.2443
3.1072 | -78.2525
3.1076 | -78.2585
3.1079 | -78.2648
3.1082 | -78.2687
3.1085 | -78.2716
3.1089 | -78.2753
3.1092 | -78.2795
3.1095 | -78.2838
3.1098 | -78.2883
3.1101 | -78.296
3.1105 | -78.3002
3.1108 | -78.3051
3.1111 | -78.3091
3.1114 | -78.3137
3.1117 | -78.3229
3.1121 | -78.3236
3.1124 | -78.3254
3.1127 | -78.3309
3.113 | -78.3361
3.1134 | -78.3397
3.1137 | -78.3425
3.114 | -78.3445
3.1143 | -78.3477
3.1146 | -78.3509
3.115 | -78.3545
3.1153 | -78.3582
3.1156 | -78.3625
3.1159 | -78.3655
3.1162 | -78.3684
3.1166 | -78.3716
3.1169 | -78.3749
3.1172 | -78.3769
3.1175 | -78.3792
3.1178 | -78.3827
3.1182 | -78.385
3.1185 | -78.3872
3.1188 | -78.3897
3.1191 | -78.3909
3.1195 | -78.3923
3.1198 | -78.3951
3.1201 | -78.3985
3.1204 | -78.4015
3.1207 | -78.4109
3.1211 | -78.4197
3.1214 | -78.4232
3.1217 | -78.4255
3.122 | -78.4285
3.1223 | -78.4319
3.1227 | -78.4347
3.123 | -78.4367
3.1233 | -78.4398
3.1236 | -78.4447
3.124 | -78.4477
3.1243 | -78.4496
3.1246 | -78.4525
3.1249 | -78.4543
3.1252 | -78.458
3.1256 | -78.4591
3.1259 | -78.4573
3.1262 | -78.4584
3.1265 | -78.4629
3.1268 | -78.4653
3.1272 | -78.4668
3.1275 | -78.4713
3.1278 | -78.4793
3.1281 | -78.4828
3.1285 | -78.4852
3.1288 | -78.4901
3.1291 | -78.495
3.1294 | -78.4981
3.1297 | -78.5001
3.1301 | -78.5024
3.1304 | -78.5034
3.1307 | -78.5055
3.131 | -78.5073
3.1313 | -78.5087
3.1317 | -78.5099
3.132 | -78.5107
3.1323 | -78.5134
3.1326 | -78.5181
3.133 | -78.5214
3.1333 | -78.525
3.1336 | -78.5274
3.1339 | -78.5302
3.1342 | -78.5332
3.1346 | -78.5339
3.1349 | -78.5353
3.1352 | -78.5378
3.1355 | -78.5404
3.1358 | -78.5406
3.1362 | -78.5413
3.1365 | -78.545
3.1368 | -78.5488
3.1371 | -78.55
3.1375 | -78.5533
3.1378 | -78.5551
3.1381 | -78.5545
3.1384 | -78.554
3.1387 | -78.554
3.1391 | -78.5554
3.1394 | -78.5542
3.1397 | -78.5543
3.14 | -78.5559
3.1403 | -78.5573
3.1407 | -78.5588
3.141 | -78.5597
3.1413 | -78.5589
3.1416 | -78.5598
3.142 | -78.5608
3.1423 | -78.5635
3.1426 | -78.5597
3.1429 | -78.5625
3.1432 | -78.5627
3.1436 | -78.564
3.1439 | -78.5643
3.1442 | -78.5644
3.1445 | -78.5641
3.1448 | -78.5651
3.1452 | -78.5647
3.1455 | -78.5644
3.1458 | -78.5626
3.1461 | -78.5571
3.1464 | -78.5521
3.1468 | -78.5522
3.1471 | -78.5532
3.1474 | -78.5526
3.1477 | -78.5529
3.1481 | -78.548
3.1484 | -78.5441
3.1487 | -78.5438
3.149 | -78.543
3.1493 | -78.5421
3.1497 | -78.5414
3.15 | -78.538
3.1503 | -78.5284
3.1506 | -78.5262
3.1509 | -78.5253
3.1513 | -78.5242
3.1516 | -78.521
3.1519 | -78.5144
3.1522 | -78.506
3.1526 | -78.4987
3.1529 | -78.49
3.1532 | -78.4823
3.1535 | -78.4793
3.1538 | -78.4754
3.1542 | -78.4685
3.1545 | -78.4667
3.1548 | -78.4594
3.1551 | -78.4582
3.1554 | -78.4555
3.1558 | -78.4535
3.1561 | -78.4509
3.1564 | -78.4475
3.1567 | -78.4447
3.1571 | -78.4394
3.1574 | -78.4294
3.1577 | -78.4242
3.158 | -78.42
3.1583 | -78.4179
3.1587 | -78.4179
3.159 | -78.4146
3.1593 | -78.4021
3.1596 | -78.3968
3.1599 | -78.3922
3.1603 | -78.3879
3.1606 | -78.3839
3.1609 | -78.3801
3.1612 | -78.3702
3.1616 | -78.3634
3.1619 | -78.3549
3.1622 | -78.3476
3.1625 | -78.3446
3.1628 | -78.3415
3.1632 | -78.338
3.1635 | -78.3341
3.1638 | -78.3299
3.1641 | -78.3242
3.1644 | -78.3181
3.1648 | -78.3125
3.1651 | -78.304
3.1654 | -78.2988
3.1657 | -78.2938
3.166 | -78.289
3.1664 | -78.2841
3.1667 | -78.2771
3.167 | -78.2721
3.1673 | -78.2667
3.1677 | -78.2625
3.168 | -78.2591
3.1683 | -78.2547
3.1686 | -78.2504
3.1689 | -78.2452
3.1693 | -78.236
3.1696 | -78.2282
3.1699 | -78.2215
3.1702 | -78.2145
3.1705 | -78.1976
3.1709 | -78.192
3.1712 | -78.1882
3.1715 | -78.1826
3.1718 | -78.1769
3.1722 | -78.1707
3.1725 | -78.1642
3.1728 | -78.1566
3.1731 | -78.1502
3.1734 | -78.1438
3.1738 | -78.1373
3.1741 | -78.1297
3.1744 | -78.1231
3.1747 | -78.1165
3.175 | -78.1099
3.1754 | -78.1029
3.1757 | -78.0933
3.176 | -78.0873
3.1763 | -78.08
3.1767 | -78.0703
3.177 | -78.0626
3.1773 | -78.0557
3.1776 | -78.048
3.1779 | -78.0414
3.1783 | -78.0344
3.1786 | -78.0276
3.1789 | -78.0208
3.1792 | -78.014
3.1795 | -78.007
3.1799 | -77.993
3.1802 | -77.9852
3.1805 | -77.9777
3.1808 | -77.9692
3.1812 | -77.9611
3.1815 | -77.952
3.1818 | -77.9443
3.1821 | -77.938
3.1824 | -77.9304
3.1828 | -77.9227
3.1831 | -77.9154
3.1834 | -77.9066
3.1837 | -77.8985
3.184 | -77.8902
3.1844 | -77.8805
3.1847 | -77.8721
3.185 | -77.8638
3.1853 | -77.8532
3.1857 | -77.8432
3.186 | -77.8351
3.1863 | -77.8266
3.1866 | -77.8179
3.1869 | -77.8093
3.1873 | -77.8009
3.1876 | -77.7936
3.1879 | -77.7857
3.1882 | -77.7766
3.1885 | -77.7676
3.1889 | -77.7568
3.1892 | -77.7428
3.1895 | -77.7336
3.1898 | -77.7247
3.1901 | -77.7165
3.1905 | -77.7071
3.1908 | -77.6971
3.1911 | -77.6862
3.1914 | -77.6749
3.1918 | -77.6652
3.1921 | -77.6562
3.1924 | -77.6473
3.1927 | -77.6374
3.193 | -77.6272
3.1934 | -77.6174
3.1937 | -77.6081
3.194 | -77.5981
3.1943 | -77.5881
3.1946 | -77.5786
3.195 | -77.5692
3.1953 | -77.5602
3.1956 | -77.551
3.1959 | -77.5421
3.1963 | -77.5335
3.1966 | -77.5237
3.1969 | -77.5145
3.1972 | -77.5063
3.1975 | -77.4971
3.1979 | -77.4876
3.1982 | -77.4777
3.1985 | -77.4657
3.1988 | -77.4562
3.1991 | -77.446
3.1995 | -77.4335
3.1998 | -77.4202
3.2001 | -77.4087
3.2004 | -77.3986
3.2008 | -77.3884
3.2011 | -77.3778
3.2014 | -77.3672
3.2017 | -77.3561
3.202 | -77.3447
3.2024 | -77.3335
3.2027 | -77.3225
3.203 | -77.3121
3.2033 | -77.3022
3.2036 | -77.2924
3.204 | -77.2824
3.2043 | -77.2719
3.2046 | -77.2613
3.2049 | -77.2517
3.2053 | -77.243
3.2056 | -77.2347
3.2059 | -77.2261
3.2062 | -77.216
3.2065 | -77.2064
3.2069 | -77.1967
3.2072 | -77.1866
3.2075 | -77.1753
3.2078 | -77.1647
3.2081 | -77.1544
3.2085 | -77.1445
3.2088 | -77.1351
3.2091 | -77.1251
3.2094 | -77.1141
3.2098 | -77.104
3.2101 | -77.0958
3.2104 | -77.0867
3.2107 | -77.0782
3.211 | -77.0704
3.2114 | -77.0609
3.2117 | -77.0508
3.212 | -77.0409
3.2123 | -77.031
3.2126 | -77.0213
3.213 | -77.0116
3.2133 | -77.003
3.2136 | -76.9965
3.2139 | -76.9884
3.2142 | -76.9784
3.2146 | -76.9678
3.2149 | -76.9578
3.2152 | -76.9481
3.2155 | -76.9385
3.2159 | -76.9291
3.2162 | -76.919
3.2165 | -76.909
3.2168 | -76.9
3.2171 | -76.8901
3.2175 | -76.8804
3.2178 | -76.8713
3.2181 | -76.8621
3.2184 | -76.8533
3.2187 | -76.8433
3.2191 | -76.8325
3.2194 | -76.823
3.2197 | -76.8141
3.22 | -76.8052
3.2204 | -76.7957
3.2207 | -76.7858
3.221 | -76.7757
3.2213 | -76.7666
3.2216 | -76.7566
3.222 | -76.7472
3.2223 | -76.7382
3.2226 | -76.7293
3.2229 | -76.7205
3.2232 | -76.7122
3.2236 | -76.7041
3.2239 | -76.6952
3.2242 | -76.6866
3.2245 | -76.6779
3.2249 | -76.6696
3.2252 | -76.6613
3.2255 | -76.6526
3.2258 | -76.6442
3.2261 | -76.637
3.2265 | -76.6295
3.2268 | -76.6222
3.2271 | -76.6161
3.2274 | -76.6086
3.2277 | -76.6003
3.2281 | -76.5915
3.2284 | -76.5833
3.2287 | -76.5761
3.229 | -76.5678
3.2294 | -76.5594
3.2297 | -76.5513
3.23 | -76.5437
3.2303 | -76.5382
3.2306 | -76.5311
3.231 | -76.5233
3.2313 | -76.5152
3.2316 | -76.5062
3.2319 | -76.498
3.2322 | -76.4906
3.2326 | -76.4832
3.2329 | -76.4763
3.2332 | -76.4696
3.2335 | -76.4623
3.2338 | -76.4551
3.2342 | -76.448
3.2345 | -76.4413
3.2348 | -76.4354
3.2351 | -76.428
3.2355 | -76.4247
3.2358 | -76.4176
3.2361 | -76.4106
3.2364 | -76.4041
3.2367 | -76.3974
3.2371 | -76.3907
3.2374 | -76.3837
3.2377 | -76.3758
3.238 | -76.3696
3.2383 | -76.3631
3.2387 | -76.3571
3.239 | -76.3516
3.2393 | -76.3466
3.2396 | -76.343
3.24 | -76.3374
3.2403 | -76.3313
3.2406 | -76.3261
3.2409 | -76.3211
3.2412 | -76.3163
3.2416 | -76.3115
3.2419 | -76.306
3.2422 | -76.3005
3.2425 | -76.295
3.2428 | -76.2893
3.2432 | -76.2841
3.2435 | -76.2796
3.2438 | -76.2745
3.2441 | -76.2689
3.2445 | -76.2639
3.2448 | -76.2637
3.2451 | -76.2587
3.2454 | -76.254
3.2457 | -76.2495
3.2461 | -76.2449
3.2464 | -76.2395
3.2467 | -76.2344
3.247 | -76.2298
3.2473 | -76.2259
3.2477 | -76.2219
3.248 | -76.2175
3.2483 | -76.2123
3.2486 | -76.2083
3.249 | -76.2051
3.2493 | -76.2023
3.2496 | -76.1987
3.2499 | -76.1951
3.2502 | -76.1915
3.2506 | -76.1876
3.2509 | -76.1842
3.2512 | -76.1815
3.2515 | -76.1787
3.2518 | -76.1756
3.2522 | -76.1726
3.2525 | -76.1697
3.2528 | -76.1668
3.2531 | -76.164
3.2535 | -76.1614
3.2538 | -76.1588
3.2541 | -76.1616
3.2544 | -76.1591
3.2547 | -76.1566
3.2551 | -76.1555
3.2554 | -76.1549
3.2557 | -76.153
3.256 | -76.1509
3.2563 | -76.1488
3.2567 | -76.1468
3.257 | -76.1447
3.2573 | -76.1423
3.2576 | -76.1404
3.258 | -76.1385
3.2583 | -76.1374
3.2586 | -76.1364
3.2589 | -76.1351
3.2592 | -76.1337
3.2596 | -76.1324
3.2599 | -76.1316
3.2602 | -76.1311
3.2605 | -76.1301
3.2608 | -76.129
3.2612 | -76.1283
3.2615 | -76.1265
3.2618 | -76.1254
3.2621 | -76.1248
3.2624 | -76.1248
3.2628 | -76.1249
3.2631 | -76.124
3.2634 | -76.129
3.2637 | -76.1287
3.2641 | -76.1287
3.2644 | -76.1286
3.2647 | -76.1285
3.265 | -76.1289
3.2653 | -76.129
3.2657 | -76.1287
3.266 | -76.1288
3.2663 | -76.1295
3.2666 | -76.1303
3.2669 | -76.1309
3.2673 | -76.1324
3.2676 | -76.1338
3.2679 | -76.1348
3.2682 | -76.1354
3.2686 | -76.1364
3.2689 | -76.1378
3.2692 | -76.1391
3.2695 | -76.1407
3.2698 | -76.1426
3.2702 | -76.1441
3.2705 | -76.1457
3.2708 | -76.1476
3.2711 | -76.1492
3.2714 | -76.1509
3.2718 | -76.1533
3.2721 | -76.1561
3.2724 | -76.1641
3.2727 | -76.1674
3.2731 | -76.1697
3.2734 | -76.1727
3.2737 | -76.1756
3.274 | -76.1785
3.2743 | -76.1819
3.2747 | -76.185
3.275 | -76.1881
3.2753 | -76.191
3.2756 | -76.1948
3.2759 | -76.1981
3.2763 | -76.2011
3.2766 | -76.2044
3.2769 | -76.2077
3.2772 | -76.2113
3.2776 | -76.2148
3.2779 | -76.2184
3.2782 | -76.2224
3.2785 | -76.2263
3.2788 | -76.2306
3.2792 | -76.235
3.2795 | -76.2391
3.2798 | -76.2428
3.2801 | -76.2469
3.2804 | -76.2513
3.2808 | -76.2564
3.2811 | -76.2617
3.2814 | -76.2666
3.2817 | -76.2765
3.282 | -76.2817
3.2824 | -76.286
3.2827 | -76.291
3.283 | -76.2961
3.2833 | -76.3013
3.2837 | -76.3062
3.284 | -76.3113
3.2843 | -76.3166
3.2846 | -76.3216
3.2849 | -76.3272
3.2853 | -76.3343
3.2856 | -76.3411
3.2859 | -76.347
3.2862 | -76.3535
3.2865 | -76.3593
3.2869 | -76.3657
3.2872 | -76.3725
3.2875 | -76.3784
3.2878 | -76.3845
3.2882 | -76.3907
3.2885 | -76.3972
3.2888 | -76.4035
3.2891 | -76.4102
3.2894 | -76.4182
3.2898 | -76.4256
3.2901 | -76.4327
3.2904 | -76.4395
3.2907 | -76.4467
3.291 | -76.4582
3.2914 | -76.4664
3.2917 | -76.4744
3.292 | -76.4819
3.2923 | -76.4895
3.2927 | -76.4971
3.293 | -76.5043
3.2933 | -76.5115
3.2936 | -76.5182
3.2939 | -76.5251
3.2943 | -76.5329
3.2946 | -76.5407
3.2949 | -76.5482
3.2952 | -76.5561
3.2955 | -76.5642
3.2959 | -76.5719
3.2962 | -76.5801
3.2965 | -76.5888
3.2968 | -76.5977
3.2972 | -76.6062
3.2975 | -76.6144
3.2978 | -76.6231
3.2981 | -76.6314
3.2984 | -76.6395
3.2988 | -76.6477
3.2991 | -76.6553
3.2994 | -76.6638
3.2997 | -76.6723
3.3 | -76.6812
3.3004 | -76.6933
3.3007 | -76.702
3.301 | -76.7115
3.3013 | -76.7206
3.3016 | -76.7298
3.302 | -76.7387
3.3023 | -76.7485
3.3026 | -76.7587
3.3029 | -76.7689
3.3033 | -76.7793
3.3036 | -76.7879
3.3039 | -76.7966
3.3042 | -76.8066
3.3045 | -76.8162
3.3049 | -76.8261
3.3052 | -76.8359
3.3055 | -76.8457
3.3058 | -76.856
3.3061 | -76.8662
3.3065 | -76.8761
3.3068 | -76.886
3.3071 | -76.8962
3.3074 | -76.9058
3.3078 | -76.9155
3.3081 | -76.9254
3.3084 | -76.9351
3.3087 | -76.9453
3.309 | -76.9566
3.3094 | -76.968
3.3097 | -76.9804
3.31 | -76.9909
3.3103 | -77.0013
3.3106 | -77.0114
3.311 | -77.0212
3.3113 | -77.0322
3.3116 | -77.0434
3.3119 | -77.0547
3.3123 | -77.0659
3.3126 | -77.0766
3.3129 | -77.0874
3.3132 | -77.0984
3.3135 | -77.1093
3.3139 | -77.1195
3.3142 | -77.1306
3.3145 | -77.1417
3.3148 | -77.1525
3.3151 | -77.164
3.3155 | -77.176
3.3158 | -77.1873
3.3161 | -77.1987
3.3164 | -77.2109
3.3168 | -77.2226
3.3171 | -77.2344
3.3174 | -77.2465
3.3177 | -77.2578
3.318 | -77.2691
3.3184 | -77.2808
3.3187 | -77.2923
3.319 | -77.3047
3.3193 | -77.3163
3.3196 | -77.3278
3.32 | -77.34
3.3203 | -77.3515
3.3206 | -77.3637
3.3209 | -77.3769
3.3213 | -77.3903
3.3216 | -77.4026
3.3219 | -77.4147
3.3222 | -77.4276
3.3225 | -77.4401
3.3229 | -77.4527
3.3232 | -77.4651
3.3235 | -77.4774
3.3238 | -77.4897
3.3241 | -77.5021
3.3245 | -77.5143
3.3248 | -77.5268
3.3251 | -77.5384
3.3254 | -77.5501
3.3258 | -77.5626
3.3261 | -77.5756
3.3264 | -77.5881
3.3267 | -77.6006
3.327 | -77.6129
3.3274 | -77.6254
3.3277 | -77.6375
3.328 | -77.6491
3.3283 | -77.6592
3.3286 | -77.6713
3.329 | -77.6828
3.3293 | -77.6929
3.3296 | -77.7046
3.3299 | -77.7178
3.3302 | -77.7301
3.3306 | -77.7426
3.3309 | -77.7555
3.3312 | -77.7692
3.3315 | -77.7827
3.3319 | -77.7961
3.3322 | -77.8088
3.3325 | -77.8218
3.3328 | -77.8342
3.3331 | -77.846
3.3335 | -77.8583
3.3338 | -77.8711
3.3341 | -77.8839
3.3344 | -77.8964
3.3347 | -77.9086
3.3351 | -77.9204
3.3354 | -77.9327
3.3357 | -77.9452
3.336 | -77.9614
3.3364 | -77.977
3.3367 | -77.9914
3.337 | -78.0066
3.3373 | -78.0165
3.3376 | -78.0306
3.338 | -78.0444
3.3383 | -78.0588
3.3386 | -78.0709
3.3389 | -78.0828
3.3392 | -78.0952
3.3396 | -78.1074
3.3399 | -78.1197
3.3402 | -78.1315
3.3405 | -78.1443
3.3409 | -78.1564
3.3412 | -78.1688
3.3415 | -78.1813
3.3418 | -78.1945
3.3421 | -78.2081
3.3425 | -78.2205
3.3428 | -78.2332
3.3431 | -78.246
3.3434 | -78.2589
3.3437 | -78.273
3.3441 | -78.2865
3.3444 | -78.2987
3.3447 | -78.3101
3.345 | -78.3218
3.3454 | -78.3328
3.3457 | -78.3429
3.346 | -78.3543
3.3463 | -78.367
3.3466 | -78.3767
3.347 | -78.3891
3.3473 | -78.4008
3.3476 | -78.413
3.3479 | -78.4265
3.3482 | -78.4396
3.3486 | -78.4504
3.3489 | -78.4617
3.3492 | -78.4737
3.3495 | -78.4852
3.3498 | -78.4976
3.3502 | -78.5098
3.3505 | -78.5216
3.3508 | -78.5345
3.3511 | -78.55
3.3515 | -78.5625
3.3518 | -78.5724
3.3521 | -78.5836
3.3524 | -78.5959
3.3527 | -78.6086
3.3531 | -78.621
3.3534 | -78.6326
3.3537 | -78.6437
3.354 | -78.6541
3.3543 | -78.6657
3.3547 | -78.6767
3.355 | -78.6877
3.3553 | -78.6995
3.3556 | -78.7114
3.356 | -78.7181
3.3563 | -78.7299
3.3566 | -78.7414
3.3569 | -78.7532
3.3572 | -78.7641
3.3576 | -78.7767
3.3579 | -78.7891
3.3582 | -78.7993
3.3585 | -78.8095
3.3588 | -78.8199
3.3592 | -78.8302
3.3595 | -78.84
3.3598 | -78.8496
3.3601 | -78.8601
3.3605 | -78.8711
3.3608 | -78.8843
3.3611 | -78.8965
3.3614 | -78.9075
3.3617 | -78.9197
3.3621 | -78.9319
3.3624 | -78.9422
3.3627 | -78.9529
3.363 | -78.9641
3.3633 | -78.9759
3.3637 | -78.9877
3.364 | -79
3.3643 | -79.0111
3.3646 | -79.0225
3.365 | -79.0322
3.3653 | -79.0379
3.3656 | -79.051
3.3659 | -79.0614
3.3662 | -79.0713
3.3666 | -79.0825
3.3669 | -79.0907
3.3672 | -79.0996
3.3675 | -79.1101
3.3678 | -79.1186
3.3682 | -79.1273
3.3685 | -79.1362
3.3688 | -79.1483
3.3691 | -79.1595
3.3695 | -79.171
3.3698 | -79.1816
3.3701 | -79.1904
3.3704 | -79.1999
3.3707 | -79.2094
3.3711 | -79.2203
3.3714 | -79.2292
3.3717 | -79.2393
3.372 | -79.2533
3.3723 | -79.2669
3.3727 | -79.2817
3.373 | -79.2944
3.3733 | -79.3025
3.3736 | -79.311
3.3739 | -79.3194
3.3743 | -79.3276
3.3746 | -79.3318
3.3749 | -79.3451
3.3752 | -79.3573
3.3756 | -79.3653
3.3759 | -79.3743
3.3762 | -79.383
3.3765 | -79.3912
3.3768 | -79.4013
3.3772 | -79.4182
3.3775 | -79.4268
3.3778 | -79.4336
3.3781 | -79.4372
3.3784 | -79.4455
3.3788 | -79.4533
3.3791 | -79.4634
3.3794 | -79.4801
3.3797 | -79.4986
3.3801 | -79.5156
3.3804 | -79.5288
3.3807 | -79.5363
3.381 | -79.5447
3.3813 | -79.5567
3.3817 | -79.5629
3.382 | -79.5763
3.3823 | -79.5829
3.3826 | -79.5881
3.3829 | -79.5918
3.3833 | -79.5971
3.3836 | -79.6028
3.3839 | -79.6057
3.3842 | -79.6119
3.3846 | -79.6207
3.3849 | -79.6351
3.3852 | -79.6433
3.3855 | -79.65
3.3858 | -79.6549
3.3862 | -79.6602
3.3865 | -79.6671
3.3868 | -79.6834
3.3871 | -79.6905
3.3874 | -79.6956
3.3878 | -79.6986
3.3881 | -79.7028
3.3884 | -79.7075
3.3887 | -79.7126
3.3891 | -79.7209
3.3894 | -79.728
3.3897 | -79.7379
3.39 | -79.7467
3.3903 | -79.7509
3.3907 | -79.7547
3.391 | -79.7589
3.3913 | -79.7624
3.3916 | -79.7669
3.3919 | -79.7734
3.3923 | -79.7796
3.3926 | -79.7847
3.3929 | -79.7867
3.3932 | -79.7928
3.3935 | -79.7967
3.3939 | -79.8017
3.3942 | -79.8055
3.3945 | -79.8079
3.3948 | -79.8144
3.3952 | -79.8187
3.3955 | -79.821
3.3958 | -79.8231
3.3961 | -79.8261
3.3964 | -79.8295
3.3968 | -79.8332
3.3971 | -79.8366
3.3974 | -79.8444
3.3977 | -79.8488
3.398 | -79.8515
3.3984 | -79.8563
3.3987 | -79.8641
3.399 | -79.8708
3.3993 | -79.8734
3.3997 | -79.8756
3.4 | -79.8775
3.4003 | -79.879
3.4006 | -79.8812
3.4009 | -79.8851
3.4013 | -79.8884
3.4016 | -79.8904
3.4019 | -79.8917
3.4022 | -79.8912
3.4025 | -79.8918
3.4029 | -79.8957
3.4032 | -79.8988
3.4035 | -79.8998
3.4038 | -79.9013
3.4042 | -79.9039
3.4045 | -79.9098
3.4048 | -79.9132
3.4051 | -79.9161
3.4054 | -79.9213
3.4058 | -79.9238
3.4061 | -79.9259
3.4064 | -79.9288
3.4067 | -79.9306
3.407 | -79.9317
3.4074 | -79.9319
3.4077 | -79.9321
3.408 | -79.9326
3.4083 | -79.9334
3.4087 | -79.9354
3.409 | -79.9385
3.4093 | -79.9393
3.4096 | -79.9406
3.4099 | -79.943
3.4103 | -79.9453
3.4106 | -79.949
3.4109 | -79.9518
3.4112 | -79.9532
3.4115 | -79.9525
3.4119 | -79.9537
3.4122 | -79.9547
3.4125 | -79.9551
3.4128 | -79.9563
3.4132 | -79.9567
3.4135 | -79.9571
3.4138 | -79.9592
3.4141 | -79.9612
3.4144 | -79.963
3.4148 | -79.9677
3.4151 | -79.9706
3.4154 | -79.9709
3.4157 | -79.9717
3.416 | -79.9721
3.4164 | -79.9729
3.4167 | -79.9731
3.417 | -79.9717
3.4173 | -79.972
3.4177 | -79.9731
3.418 | -79.9742
3.4183 | -79.9781
3.4186 | -79.9815
3.4189 | -79.983
3.4193 | -79.9832
3.4196 | -79.9818
3.4199 | -79.9808
3.4202 | -79.9813
3.4205 | -79.9832
3.4209 | -79.9821
3.4212 | -79.9842
3.4215 | -79.9859
3.4218 | -79.9847
3.4221 | -79.9843
3.4225 | -79.9843
3.4228 | -79.9851
3.4231 | -79.987
3.4234 | -79.9881
3.4238 | -79.987
3.4241 | -79.9856
3.4244 | -79.9858
3.4247 | -79.9862
3.425 | -79.9857
3.4254 | -79.9855
3.4257 | -79.9851
3.426 | -79.9855
3.4263 | -79.9858
3.4266 | -79.986
3.427 | -79.9862
3.4273 | -79.9865
3.4276 | -79.987
3.4279 | -79.9872
3.4283 | -79.9869
3.4286 | -79.9864
3.4289 | -79.9866
3.4292 | -79.987
3.4295 | -79.9878
3.4299 | -79.9918
3.4302 | -79.992
3.4305 | -79.9973
3.4308 | -79.9997
3.4311 | -80.0005
3.4315 | -80.0018
3.4318 | -80.0041
3.4321 | -80.0039
3.4324 | -80.0051
3.4328 | -80.0071
3.4331 | -80.0086
3.4334 | -80.009
3.4337 | -80.0084
3.434 | -80.0078
3.4344 | -80.0063
3.4347 | -80.0048
3.435 | -80.0044
3.4353 | -80.0046
3.4356 | -80.0038
3.436 | -80.0023
3.4363 | -80.0009
3.4366 | -80.0011
3.4369 | -80.0027
3.4373 | -80.0023
3.4376 | -80.002
3.4379 | -80.0013
3.4382 | -80.0009
3.4385 | -80.0019
3.4389 | -80.0022
3.4392 | -80.0024
3.4395 | -80.0003
3.4398 | -80.0002
3.4401 | -80
3.4405 | -80.0002
3.4408 | -80.0007
3.4411 | -79.9989
3.4414 | -79.9977
3.4417 | -79.9959
3.4421 | -79.9947
3.4424 | -79.9944
3.4427 | -79.9933
3.443 | -79.9913
3.4434 | -79.9906
3.4437 | -79.9899
3.444 | -79.9892
3.4443 | -79.9879
3.4446 | -79.988
3.445 | -79.9888
3.4453 | -79.9876
3.4456 | -79.9869
3.4459 | -79.9852
3.4462 | -79.9846
3.4466 | -79.9838
3.4469 | -79.9832
3.4472 | -79.9813
3.4475 | -79.9821
3.4479 | -79.98
3.4482 | -79.9785
3.4485 | -79.9782
3.4488 | -79.9779
3.4491 | -79.9771
3.4495 | -79.9757
3.4498 | -79.9762
3.4501 | -79.9773
3.4504 | -79.9745
3.4507 | -79.9704
3.4511 | -79.9696
3.4514 | -79.969
3.4517 | -79.968
3.452 | -79.9664
3.4524 | -79.9667
3.4527 | -79.9667
3.453 | -79.9676
3.4533 | -79.9669
3.4536 | -79.9663
3.454 | -79.9656
3.4543 | -79.9647
3.4546 | -79.9633
3.4549 | -79.9618
3.4552 | -79.9604
3.4556 | -79.9589
3.4559 | -79.9581
3.4562 | -79.957
3.4565 | -79.9557
3.4569 | -79.9534
3.4572 | -79.9527
3.4575 | -79.9505
3.4578 | -79.9503
3.4581 | -79.9498
3.4585 | -79.9493
3.4588 | -79.9491
3.4591 | -79.948
3.4594 | -79.9472
3.4597 | -79.9463
3.4601 | -79.9455
3.4604 | -79.9445
3.4607 | -79.9453
3.461 | -79.944
3.4614 | -79.942
3.4617 | -79.9393
3.462 | -79.935
3.4623 | -79.9317
3.4626 | -79.9284
3.463 | -79.9257
3.4633 | -79.923
3.4636 | -79.9176
3.4639 | -79.914
3.4642 | -79.9171
3.4646 | -79.9171
3.4649 | -79.9151
3.4652 | -79.912
3.4655 | -79.9085
3.4658 | -79.9046
3.4662 | -79.9008
3.4665 | -79.8966
3.4668 | -79.8934
3.4671 | -79.8882
3.4675 | -79.8838
3.4678 | -79.8818
3.4681 | -79.8807
3.4684 | -79.879
3.4687 | -79.8779
3.4691 | -79.8761
3.4694 | -79.8737
3.4697 | -79.8701
3.47 | -79.8658
3.4703 | -79.8619
3.4707 | -79.8613
3.471 | -79.8614
3.4713 | -79.8582
3.4716 | -79.855
3.472 | -79.8498
3.4723 | -79.8464
3.4726 | -79.8436
3.4729 | -79.8395
3.4732 | -79.8354
3.4736 | -79.8316
3.4739 | -79.8293
3.4742 | -79.8265
3.4745 | -79.8217
3.4748 | -79.8183
3.4752 | -79.8141
3.4755 | -79.8102
3.4758 | -79.8077
3.4761 | -79.8066
3.4765 | -79.8038
3.4768 | -79.8014
3.4771 | -79.7968
3.4774 | -79.7894
3.4777 | -79.7848
3.4781 | -79.7831
3.4784 | -79.7809
3.4787 | -79.7791
3.479 | -79.7736
3.4793 | -79.7663
3.4797 | -79.7622
3.48 | -79.7591
3.4803 | -79.7575
3.4806 | -79.7558
3.481 | -79.7524
3.4813 | -79.7485
3.4816 | -79.7459
3.4819 | -79.7435
3.4822 | -79.7407
3.4826 | -79.7373
3.4829 | -79.7345
3.4832 | -79.7312
3.4835 | -79.7266
3.4838 | -79.7219
3.4842 | -79.7188
3.4845 | -79.7156
3.4848 | -79.7119
3.4851 | -79.7083
3.4855 | -79.7041
3.4858 | -79.6997
3.4861 | -79.6965
3.4864 | -79.6924
3.4867 | -79.6886
3.4871 | -79.6857
3.4874 | -79.6836
3.4877 | -79.6798
3.488 | -79.6754
3.4883 | -79.6719
3.4887 | -79.669
3.489 | -79.6647
3.4893 | -79.6625
3.4896 | -79.6593
3.4899 | -79.6563
3.4903 | -79.6537
3.4906 | -79.6492
3.4909 | -79.6431
3.4912 | -79.6375
3.4916 | -79.6342
3.4919 | -79.632
3.4922 | -79.6289
3.4925 | -79.6259
3.4928 | -79.6214
3.4932 | -79.6164
3.4935 | -79.6121
3.4938 | -79.6085
3.4941 | -79.6029
3.4944 | -79.5986
3.4948 | -79.5936
3.4951 | -79.5877
3.4954 | -79.5838
3.4957 | -79.5804
3.4961 | -79.5765
3.4964 | -79.5741
3.4967 | -79.5704
3.497 | -79.5669
3.4973 | -79.562
3.4977 | -79.5581
3.498 | -79.5548
3.4983 | -79.551
3.4986 | -79.5467
3.4989 | -79.5426
3.4993 | -79.5385
3.4996 | -79.5344
3.4999 | -79.5309
3.5002 | -79.525
3.5006 | -79.5203
3.5009 | -79.5174
3.5012 | -79.5136
3.5015 | -79.508
3.5018 | -79.5034
3.5022 | -79.4991
3.5025 | -79.4947
3.5028 | -79.4898
3.5031 | -79.4855
3.5034 | -79.4826
3.5038 | -79.479
3.5041 | -79.4744
3.5044 | -79.4687
3.5047 | -79.4646
3.5051 | -79.4607
3.5054 | -79.4562
3.5057 | -79.4519
3.506 | -79.4469
3.5063 | -79.4395
3.5067 | -79.4322
3.507 | -79.4273
3.5073 | -79.422
3.5076 | -79.4155
3.5079 | -79.4105
3.5083 | -79.4057
3.5086 | -79.4004
3.5089 | -79.3953
3.5092 | -79.3908
3.5095 | -79.3862
3.5099 | -79.3811
3.5102 | -79.3761
3.5105 | -79.3709
3.5108 | -79.365
3.5112 | -79.3606
3.5115 | -79.3557
3.5118 | -79.3511
3.5121 | -79.3459
3.5124 | -79.3409
3.5128 | -79.3351
3.5131 | -79.3299
3.5134 | -79.3247
3.5137 | -79.3192
3.514 | -79.3143
3.5144 | -79.31
3.5147 | -79.3052
3.515 | -79.3004
3.5153 | -79.2954
3.5157 | -79.2908
3.516 | -79.2867
3.5163 | -79.2817
3.5166 | -79.2765
3.5169 | -79.2698
3.5173 | -79.263
3.5176 | -79.2579
3.5179 | -79.253
3.5182 | -79.2471
3.5185 | -79.2423
3.5189 | -79.2377
3.5192 | -79.2329
3.5195 | -79.2272
3.5198 | -79.2221
3.5202 | -79.2176
3.5205 | -79.213
3.5208 | -79.2078
3.5211 | -79.2034
3.5214 | -79.1998
3.5218 | -79.1951
3.5221 | -79.1898
3.5224 | -79.1851
3.5227 | -79.1799
3.523 | -79.175
3.5234 | -79.1695
3.5237 | -79.1649
3.524 | -79.1591
3.5243 | -79.1531
3.5247 | -79.1473
3.525 | -79.1422
3.5253 | -79.1368
3.5256 | -79.1324
3.5259 | -79.1285
3.5263 | -79.1232
3.5266 | -79.1175
3.5269 | -79.1126
3.5272 | -79.1059
3.5275 | -79.1001
3.5279 | -79.0949
3.5282 | -79.0897
3.5285 | -79.0834
3.5288 | -79.0777
3.5292 | -79.0728
3.5295 | -79.0678
3.5298 | -79.0626
3.5301 | -79.0578
3.5304 | -79.0524
3.5308 | -79.0477
3.5311 | -79.0423
3.5314 | -79.037
3.5317 | -79.0318
3.532 | -79.0262
3.5324 | -79.0215
3.5327 | -79.0166
3.533 | -79.0117
3.5333 | -79.0061
3.5337 | -78.9995
3.534 | -78.9941
3.5343 | -78.9888
3.5346 | -78.9833
3.5349 | -78.9777
3.5353 | -78.9721
3.5356 | -78.9676
3.5359 | -78.9627
3.5362 | -78.9579
3.5365 | -78.954
3.5369 | -78.9474
3.5372 | -78.942
3.5375 | -78.9368
3.5378 | -78.9317
3.5381 | -78.9261
3.5385 | -78.921
3.5388 | -78.9154
3.5391 | -78.9099
3.5394 | -78.904
3.5398 | -78.899
3.5401 | -78.8935
3.5404 | -78.8885
3.5407 | -78.8839
3.541 | -78.8787
3.5414 | -78.8724
3.5417 | -78.8654
3.542 | -78.8619
3.5423 | -78.857
3.5426 | -78.8517
3.543 | -78.8465
3.5433 | -78.8414
3.5436 | -78.8363
3.5439 | -78.831
3.5443 | -78.8258
3.5446 | -78.821
3.5449 | -78.816
3.5452 | -78.8111
3.5455 | -78.8072
3.5459 | -78.803
3.5462 | -78.7988
3.5465 | -78.7947
3.5468 | -78.7914
3.5471 | -78.7878
3.5475 | -78.7832
3.5478 | -78.7776
3.5481 | -78.7711
3.5484 | -78.764
3.5488 | -78.7585
3.5491 | -78.754
3.5494 | -78.7503
3.5497 | -78.7459
3.55 | -78.7412
3.5504 | -78.7367
3.5507 | -78.7335
3.551 | -78.7287
3.5513 | -78.7247
3.5516 | -78.7207
3.552 | -78.7162
3.5523 | -78.7114
3.5526 | -78.7075
3.5529 | -78.7031
3.5533 | -78.6989
3.5536 | -78.6939
3.5539 | -78.6889
3.5542 | -78.6845
3.5545 | -78.6805
3.5549 | -78.6761
3.5552 | -78.6713
3.5555 | -78.6671
3.5558 | -78.6623
3.5561 | -78.658
3.5565 | -78.654
3.5568 | -78.6503
3.5571 | -78.6453
3.5574 | -78.6398
3.5577 | -78.6349
3.5581 | -78.63
3.5584 | -78.6258
3.5587 | -78.6221
3.559 | -78.6175
3.5594 | -78.6149
3.5597 | -78.6126
3.56 | -78.6101
3.5603 | -78.6057
3.5606 | -78.6015
3.561 | -78.5983
3.5613 | -78.5973
3.5616 | -78.5948
3.5619 | -78.5908
3.5622 | -78.5866
3.5626 | -78.5832
3.5629 | -78.5797
3.5632 | -78.5763
3.5635 | -78.5726
3.5639 | -78.5691
3.5642 | -78.5649
3.5645 | -78.5612
3.5648 | -78.559
3.5651 | -78.5555
3.5655 | -78.5519
3.5658 | -78.549
3.5661 | -78.5459
3.5664 | -78.5424
3.5667 | -78.5389
3.5671 | -78.5357
3.5674 | -78.5332
3.5677 | -78.5312
3.568 | -78.5281
3.5684 | -78.5248
3.5687 | -78.522
3.569 | -78.5197
3.5693 | -78.519
3.5696 | -78.5174
3.57 | -78.5139
3.5703 | -78.5112
3.5706 | -78.5082
3.5709 | -78.5053
3.5712 | -78.5024
3.5716 | -78.4997
3.5719 | -78.4965
3.5722 | -78.4936
3.5725 | -78.4909
3.5729 | -78.4885
3.5732 | -78.4848
3.5735 | -78.4817
3.5738 | -78.4793
3.5741 | -78.4768
3.5745 | -78.4742
3.5748 | -78.4721
3.5751 | -78.4706
3.5754 | -78.4686
3.5757 | -78.4666
3.5761 | -78.4645
3.5764 | -78.4623
3.5767 | -78.4605
3.577 | -78.4583
3.5774 | -78.4565
3.5777 | -78.4554
3.578 | -78.4537
3.5783 | -78.4523
3.5786 | -78.4513
3.579 | -78.4479
3.5793 | -78.4467
3.5796 | -78.4449
3.5799 | -78.4423
3.5802 | -78.4403
3.5806 | -78.4387
3.5809 | -78.4365
3.5812 | -78.435
3.5815 | -78.4339
3.5818 | -78.4325
3.5822 | -78.4312
3.5825 | -78.43
3.5828 | -78.4289
3.5831 | -78.4281
3.5835 | -78.4268
3.5838 | -78.4259
3.5841 | -78.425
3.5844 | -78.424
3.5847 | -78.4233
3.5851 | -78.4222
3.5854 | -78.4218
3.5857 | -78.4217
3.586 | -78.4221
3.5863 | -78.4232
3.5867 | -78.4229
3.587 | -78.422
3.5873 | -78.421
3.5876 | -78.423
3.588 | -78.4217
3.5883 | -78.421
3.5886 | -78.4205
3.5889 | -78.4198
3.5892 | -78.4197
3.5896 | -78.4204
3.5899 | -78.4212
3.5902 | -78.4208
3.5905 | -78.4201
3.5908 | -78.4204
3.5912 | -78.4192
3.5915 | -78.4191
3.5918 | -78.4193
3.5921 | -78.4194
3.5925 | -78.4191
3.5928 | -78.4198
3.5931 | -78.4207
3.5934 | -78.4217
3.5937 | -78.4225
3.5941 | -78.4218
3.5944 | -78.4212
3.5947 | -78.421
3.595 | -78.4216
3.5953 | -78.4224
3.5957 | -78.4222
3.596 | -78.4222
3.5963 | -78.4223
3.5966 | -78.4229
3.597 | -78.4272
3.5973 | -78.4274
3.5976 | -78.4277
3.5979 | -78.428
3.5982 | -78.4285
3.5986 | -78.4293
3.5989 | -78.4303
3.5992 | -78.4315
3.5995 | -78.432
3.5998 | -78.4328
3.6002 | -78.4334
3.6005 | -78.4346
3.6008 | -78.4362
3.6011 | -78.4377
3.6014 | -78.4392
3.6018 | -78.4413
3.6021 | -78.4431
3.6024 | -78.445
3.6027 | -78.4463
3.6031 | -78.4478
3.6034 | -78.4492
3.6037 | -78.4504
3.604 | -78.4519
3.6043 | -78.4537
3.6047 | -78.4559
3.605 | -78.4582
3.6053 | -78.4599
3.6056 | -78.4616
3.6059 | -78.4638
3.6063 | -78.4692
3.6066 | -78.472
3.6069 | -78.4744
3.6072 | -78.4752
3.6076 | -78.4768
3.6079 | -78.4792
3.6082 | -78.4806
3.6085 | -78.4828
3.6088 | -78.4855
3.6092 | -78.488
3.6095 | -78.4915
3.6098 | -78.4946
3.6101 | -78.4972
3.6104 | -78.5006
3.6108 | -78.5039
3.6111 | -78.5061
3.6114 | -78.5085
3.6117 | -78.511
3.6121 | -78.5121
3.6124 | -78.5127
3.6127 | -78.5153
3.613 | -78.5178
3.6133 | -78.5204
3.6137 | -78.5233
3.614 | -78.5257
3.6143 | -78.5279
3.6146 | -78.5309
3.6149 | -78.5336
3.6153 | -78.5366
3.6156 | -78.5425
3.6159 | -78.5461
3.6162 | -78.5493
3.6166 | -78.5524
3.6169 | -78.5556
3.6172 | -78.558
3.6175 | -78.5606
3.6178 | -78.5643
3.6182 | -78.5682
3.6185 | -78.572
3.6188 | -78.575
3.6191 | -78.5772
3.6194 | -78.5801
3.6198 | -78.5838
3.6201 | -78.5878
3.6204 | -78.5918
3.6207 | -78.5954
3.6211 | -78.599
3.6214 | -78.6058
3.6217 | -78.6107
3.622 | -78.6153
3.6223 | -78.6186
3.6227 | -78.6223
3.623 | -78.6263
3.6233 | -78.6301
3.6236 | -78.634
3.6239 | -78.6373
3.6243 | -78.6412
3.6246 | -78.644
3.6249 | -78.6468
3.6252 | -78.651
3.6255 | -78.6547
3.6259 | -78.6586
3.6262 | -78.6635
3.6265 | -78.6684
3.6268 | -78.6707
3.6272 | -78.6742
3.6275 | -78.6774
3.6278 | -78.6789
3.6281 | -78.6825
3.6284 | -78.6866
3.6288 | -78.6906
3.6291 | -78.6943
3.6294 | -78.6979
3.6297 | -78.7013
3.63 | -78.7052
3.6304 | -78.7096
3.6307 | -78.7138
3.631 | -78.7178
3.6313 | -78.7225
3.6317 | -78.7286
3.632 | -78.7332
3.6323 | -78.7372
3.6326 | -78.7415
3.6329 | -78.7455
3.6333 | -78.7506
3.6336 | -78.7547
3.6339 | -78.7588
3.6342 | -78.7649
3.6345 | -78.7696
3.6349 | -78.7741
3.6352 | -78.7783
3.6355 | -78.7829
3.6358 | -78.787
3.6362 | -78.7915
3.6365 | -78.7961
3.6368 | -78.8009
3.6371 | -78.8056
3.6374 | -78.8105
3.6378 | -78.8175
3.6381 | -78.8239
3.6384 | -78.8269
3.6387 | -78.8314
3.639 | -78.8368
3.6394 | -78.8419
3.6397 | -78.8475
3.64 | -78.8533
3.6403 | -78.8585
3.6407 | -78.8634
3.641 | -78.8681
3.6413 | -78.8731
3.6416 | -78.8776
3.6419 | -78.8818
3.6423 | -78.888
3.6426 | -78.8941
3.6429 | -78.8981
3.6432 | -78.9031
3.6435 | -78.9091
3.6439 | -78.9146
3.6442 | -78.9203
3.6445 | -78.9251
3.6448 | -78.9303
3.6452 | -78.9351
3.6455 | -78.9397
3.6458 | -78.9454
3.6461 | -78.9519
3.6464 | -78.957
3.6468 | -78.962
3.6471 | -78.9675
3.6474 | -78.973
3.6477 | -78.9791
3.648 | -78.9843
3.6484 | -78.9893
3.6487 | -78.9945
3.649 | -78.9996
3.6493 | -79.0048
3.6496 | -79.01
3.65 | -79.014
3.6503 | -79.0179
3.6506 | -79.0229
3.6509 | -79.0278
3.6513 | -79.033
3.6516 | -79.0384
3.6519 | -79.0437
3.6522 | -79.049
3.6525 | -79.0548
3.6529 | -79.0605
3.6532 | -79.0669
3.6535 | -79.0728
3.6538 | -79.0788
3.6541 | -79.0852
3.6545 | -79.0902
3.6548 | -79.0962
3.6551 | -79.1014
3.6554 | -79.1068
3.6558 | -79.1123
3.6561 | -79.1174
3.6564 | -79.122
3.6567 | -79.1272
3.657 | -79.1327
3.6574 | -79.1383
3.6577 | -79.1432
3.658 | -79.1482
3.6583 | -79.1531
3.6586 | -79.1589
3.659 | -79.164
3.6593 | -79.1693
3.6596 | -79.1742
3.6599 | -79.1795
3.6603 | -79.1848
3.6606 | -79.1906
3.6609 | -79.1966
3.6612 | -79.2023
3.6615 | -79.2073
3.6619 | -79.2125
3.6622 | -79.2174
3.6625 | -79.2225
3.6628 | -79.2274
3.6631 | -79.2336
3.6635 | -79.2393
3.6638 | -79.2435
3.6641 | -79.2481
3.6644 | -79.253
3.6648 | -79.2584
3.6651 | -79.2634
3.6654 | -79.2686
3.6657 | -79.2756
3.666 | -79.282
3.6664 | -79.2875
3.6667 | -79.2927
3.667 | -79.2987
3.6673 | -79.3038
3.6676 | -79.3089
3.668 | -79.3156
3.6683 | -79.3221
3.6686 | -79.3275
3.6689 | -79.3316
3.6692 | -79.3349
3.6696 | -79.3401
3.6699 | -79.3451
3.6702 | -79.3506
3.6705 | -79.3581
3.6709 | -79.3648
3.6712 | -79.3691
3.6715 | -79.3725
3.6718 | -79.3765
3.6721 | -79.3815
3.6725 | -79.3866
3.6728 | -79.3913
3.6731 | -79.396
3.6734 | -79.4009
3.6737 | -79.4057
3.6741 | -79.4097
3.6744 | -79.4146
3.6747 | -79.4199
3.675 | -79.4255
3.6754 | -79.4312
3.6757 | -79.4362
3.676 | -79.4395
3.6763 | -79.4438
3.6766 | -79.4485
3.677 | -79.4535
3.6773 | -79.4583
3.6776 | -79.4632
3.6779 | -79.4683
3.6782 | -79.4732
3.6786 | -79.4781
3.6789 | -79.4832
3.6792 | -79.4887
3.6795 | -79.4939
3.6799 | -79.499
3.6802 | -79.5041
3.6805 | -79.5078
3.6808 | -79.5125
3.6811 | -79.5173
3.6815 | -79.5218
3.6818 | -79.5269
3.6821 | -79.5324
3.6824 | -79.5377
3.6827 | -79.5423
3.6831 | -79.5472
3.6834 | -79.5527
3.6837 | -79.5576
3.684 | -79.5612
3.6844 | -79.5647
3.6847 | -79.5681
3.685 | -79.5724
3.6853 | -79.5768
3.6856 | -79.5807
3.686 | -79.5853
3.6863 | -79.5883
3.6866 | -79.5902
3.6869 | -79.5936
3.6872 | -79.5974
3.6876 | -79.601
3.6879 | -79.605
3.6882 | -79.6091
3.6885 | -79.6146
3.6889 | -79.6211
3.6892 | -79.6258
3.6895 | -79.631
3.6898 | -79.6325
3.6901 | -79.6367
3.6905 | -79.6407
3.6908 | -79.6453
3.6911 | -79.6505
3.6914 | -79.6546
3.6917 | -79.6587
3.6921 | -79.6626
3.6924 | -79.6669
3.6927 | -79.6711
3.693 | -79.6751
3.6934 | -79.6796
3.6937 | -79.6837
3.694 | -79.6896
3.6943 | -79.694
3.6946 | -79.6978
3.695 | -79.7019
3.6953 | -79.7073
3.6956 | -79.7112
3.6959 | -79.7172
3.6962 | -79.7218
3.6966 | -79.7261
3.6969 | -79.7299
3.6972 | -79.7339
3.6975 | -79.7382
3.6978 | -79.7425
3.6982 | -79.7461
3.6985 | -79.7499
3.6988 | -79.7553
3.6991 | -79.7578
3.6995 | -79.7659
3.6998 | -79.7718
3.7001 | -79.7746
3.7004 | -79.7775
3.7007 | -79.7814
3.7011 | -79.7853
3.7014 | -79.7906
3.7017 | -79.7941
3.702 | -79.7972
3.7023 | -79.8004
3.7027 | -79.8033
3.703 | -79.8065
3.7033 | -79.8104
3.7036 | -79.8137
3.704 | -79.8173
3.7043 | -79.8205
3.7046 | -79.8218
3.7049 | -79.8232
3.7052 | -79.8268
3.7056 | -79.8311
3.7059 | -79.834
3.7062 | -79.8382
3.7065 | -79.8416
3.7068 | -79.8437
3.7072 | -79.8455
3.7075 | -79.8487
3.7078 | -79.8527
3.7081 | -79.8552
3.7085 | -79.8583
3.7088 | -79.861
3.7091 | -79.8635
3.7094 | -79.8666
3.7097 | -79.8706
3.7101 | -79.8729
3.7104 | -79.8748
3.7107 | -79.8788
3.711 | -79.8838
3.7113 | -79.8895
3.7117 | -79.8943
3.712 | -79.8995
3.7123 | -79.9046
3.7126 | -79.9088
3.713 | -79.9157
3.7133 | -79.9214
3.7136 | -79.9234
3.7139 | -79.9262
3.7142 | -79.9289
3.7146 | -79.9325
3.7149 | -79.937
3.7152 | -79.942
3.7155 | -79.946
3.7158 | -79.9489
3.7162 | -79.9516
3.7165 | -79.9544
3.7168 | -79.9599
3.7171 | -79.9653
3.7174 | -79.9669
3.7178 | -79.9687
3.7181 | -79.9705
3.7184 | -79.9708
3.7187 | -79.9709
3.7191 | -79.9738
3.7194 | -79.9756
3.7197 | -79.9786
3.72 | -79.9829
3.7203 | -79.986
3.7207 | -79.9878
3.721 | -79.9911
3.7213 | -79.9945
3.7216 | -79.9977
3.7219 | -80.0027
3.7223 | -80.0049
3.7226 | -80.0078
3.7229 | -80.0117
3.7232 | -80.0147
3.7236 | -80.0184
3.7239 | -80.0211
3.7242 | -80.0238
3.7245 | -80.0282
3.7248 | -80.0305
3.7252 | -80.0339
3.7255 | -80.0371
3.7258 | -80.0385
3.7261 | -80.04
3.7264 | -80.0411
3.7268 | -80.0426
3.7271 | -80.044
3.7274 | -80.0461
3.7277 | -80.0498
3.7281 | -80.0572
3.7284 | -80.0623
3.7287 | -80.0636
3.729 | -80.0645
3.7293 | -80.066
3.7297 | -80.0709
3.73 | -80.0771
3.7303 | -80.0802
3.7306 | -80.0823
3.7309 | -80.0846
3.7313 | -80.0866
3.7316 | -80.0876
3.7319 | -80.0894
3.7322 | -80.0906
3.7326 | -80.093
3.7329 | -80.097
3.7332 | -80.0998
3.7335 | -80.1015
3.7338 | -80.1023
3.7342 | -80.105
3.7345 | -80.1076
3.7348 | -80.1094
3.7351 | -80.1119
3.7354 | -80.1138
3.7358 | -80.1154
3.7361 | -80.1174
3.7364 | -80.12
3.7367 | -80.1217
3.737 | -80.1227
3.7374 | -80.1238
3.7377 | -80.1256
3.738 | -80.1256
3.7383 | -80.1256
3.7387 | -80.1268
3.739 | -80.128
3.7393 | -80.1305
3.7396 | -80.1316
3.7399 | -80.1318
3.7403 | -80.1325
3.7406 | -80.1327
3.7409 | -80.1341
3.7412 | -80.1363
3.7415 | -80.1375
3.7419 | -80.1398
3.7422 | -80.1435
3.7425 | -80.1467
3.7428 | -80.148
3.7432 | -80.1497
3.7435 | -80.1508
3.7438 | -80.1508
3.7441 | -80.1529
3.7444 | -80.1543
3.7448 | -80.155
3.7451 | -80.1564
3.7454 | -80.1588
3.7457 | -80.1625
3.746 | -80.1634
3.7464 | -80.1647
3.7467 | -80.1658
3.747 | -80.1676
3.7473 | -80.1687
3.7477 | -80.1694
3.748 | -80.1706
3.7483 | -80.1719
3.7486 | -80.1733
3.7489 | -80.1733
3.7493 | -80.1744
3.7496 | -80.1755
3.7499 | -80.1778
3.7502 | -80.1796
3.7505 | -80.1811
3.7509 | -80.1821
3.7512 | -80.1834
3.7515 | -80.1851
3.7518 | -80.1861
3.7522 | -80.1887
3.7525 | -80.1907
3.7528 | -80.1923
3.7531 | -80.1942
3.7534 | -80.1963
3.7538 | -80.1965
3.7541 | -80.1983
3.7544 | -80.1987
3.7547 | -80.1985
3.755 | -80.1985
3.7554 | -80.1991
3.7557 | -80.1996
3.756 | -80.2005
3.7563 | -80.2021
3.7567 | -80.2037
3.757 | -80.2053
3.7573 | -80.2065
3.7576 | -80.2067
3.7579 | -80.2075
3.7583 | -80.2086
3.7586 | -80.2105
3.7589 | -80.2143
3.7592 | -80.2176
3.7595 | -80.2171
3.7599 | -80.2185
3.7602 | -80.2212
3.7605 | -80.222
3.7608 | -80.2219
3.7612 | -80.223
3.7615 | -80.2251
3.7618 | -80.2263
3.7621 | -80.2266
3.7624 | -80.2263
3.7628 | -80.227
3.7631 | -80.2286
3.7634 | -80.2303
3.7637 | -80.2309
3.764 | -80.2322
3.7644 | -80.2333
3.7647 | -80.2348
3.765 | -80.2371
3.7653 | -80.2366
3.7656 | -80.2356
3.766 | -80.2369
3.7663 | -80.2384
3.7666 | -80.2392
3.7669 | -80.2403
3.7673 | -80.2419
3.7676 | -80.2428
3.7679 | -80.2437
3.7682 | -80.2438
3.7685 | -80.2449
3.7689 | -80.2454
3.7692 | -80.2468
3.7695 | -80.248
3.7698 | -80.2481
3.7701 | -80.2501
3.7705 | -80.2526
3.7708 | -80.2547
3.7711 | -80.2556
3.7714 | -80.2552
3.7718 | -80.2556
3.7721 | -80.2573
3.7724 | -80.2583
3.7727 | -80.26
3.773 | -80.262
3.7734 | -80.2626
3.7737 | -80.2627
3.774 | -80.2624
3.7743 | -80.2628
3.7746 | -80.2637
3.775 | -80.2645
3.7753 | -80.2652
3.7756 | -80.2663
3.7759 | -80.2673
3.7763 | -80.2683
3.7766 | -80.269
3.7769 | -80.2698
3.7772 | -80.2716
3.7775 | -80.2728
3.7779 | -80.275
3.7782 | -80.2773
3.7785 | -80.2797
3.7788 | -80.2805
3.7791 | -80.2815
3.7795 | -80.2824
3.7798 | -80.2838
3.7801 | -80.2838
3.7804 | -80.285
3.7808 | -80.2856
3.7811 | -80.2891
3.7814 | -80.2916
3.7817 | -80.2922
3.782 | -80.2929
3.7824 | -80.2931
3.7827 | -80.2955
3.783 | -80.2969
3.7833 | -80.2974
3.7836 | -80.298
3.784 | -80.2999
3.7843 | -80.3007
3.7846 | -80.3024
3.7849 | -80.3042
3.7853 | -80.3048
3.7856 | -80.3045
3.7859 | -80.3052
3.7862 | -80.3066
3.7865 | -80.3083
3.7869 | -80.3087
3.7872 | -80.3084
3.7875 | -80.3084
3.7878 | -80.3094
3.7881 | -80.3117
3.7885 | -80.3134
3.7888 | -80.3142
3.7891 | -80.3153
3.7894 | -80.3169
3.7897 | -80.3172
3.7901 | -80.317
3.7904 | -80.3176
3.7907 | -80.317
3.791 | -80.3179
3.7914 | -80.3201
3.7917 | -80.3206
3.792 | -80.3204
3.7923 | -80.3212
3.7926 | -80.322
3.793 | -80.3244
3.7933 | -80.3252
3.7936 | -80.3266
3.7939 | -80.3283
3.7942 | -80.3306
3.7946 | -80.332
3.7949 | -80.3333
3.7952 | -80.3348
3.7955 | -80.3356
3.7959 | -80.3362
3.7962 | -80.3365
3.7965 | -80.3405
3.7968 | -80.3459
3.7971 | -80.3478
3.7975 | -80.3484
3.7978 | -80.3493
3.7981 | -80.35
3.7984 | -80.3508
3.7987 | -80.3513
3.7991 | -80.3518
3.7994 | -80.3531
3.7997 | -80.3529
3.8 | -80.354
3.8004 | -80.3548
3.8007 | -80.3556
3.801 | -80.3564
3.8013 | -80.357
3.8016 | -80.3573
3.802 | -80.3585
3.8023 | -80.359
3.8026 | -80.3598
3.8029 | -80.3608
3.8032 | -80.3625
3.8036 | -80.365
3.8039 | -80.3681
3.8042 | -80.3695
3.8045 | -80.3708
3.8049 | -80.3712
3.8052 | -80.3702
3.8055 | -80.3704
3.8058 | -80.3707
3.8061 | -80.3716
3.8065 | -80.3723
3.8068 | -80.3732
3.8071 | -80.3742
3.8074 | -80.375
3.8077 | -80.3754
3.8081 | -80.377
3.8084 | -80.3787
3.8087 | -80.3771
3.809 | -80.3762
3.8093 | -80.3776
3.8097 | -80.379
3.81 | -80.3793
3.8103 | -80.3796
3.8106 | -80.3802
3.811 | -80.3811
3.8113 | -80.3826
3.8116 | -80.3834
3.8119 | -80.3845
3.8122 | -80.3856
3.8126 | -80.3863
3.8129 | -80.387
3.8132 | -80.3888
3.8135 | -80.39
3.8138 | -80.3908
3.8142 | -80.3918
3.8145 | -80.3925
3.8148 | -80.3915
3.8151 | -80.3924
3.8155 | -80.3933
3.8158 | -80.3938
3.8161 | -80.3946
3.8164 | -80.3962
3.8167 | -80.3976
3.8171 | -80.3981
3.8174 | -80.3992
3.8177 | -80.4009
3.818 | -80.4015
3.8183 | -80.4026
3.8187 | -80.4038
3.819 | -80.4039
3.8193 | -80.4036
3.8196 | -80.4033
3.82 | -80.404
3.8203 | -80.4045
3.8206 | -80.4038
3.8209 | -80.404
3.8212 | -80.4058
3.8216 | -80.4065
3.8219 | -80.4071
3.8222 | -80.4068
3.8225 | -80.4071
3.8228 | -80.4076
3.8232 | -80.4078
3.8235 | -80.4078
3.8238 | -80.4072
3.8241 | -80.4075
3.8245 | -80.4082
3.8248 | -80.4087
3.8251 | -80.4093
3.8254 | -80.4108
3.8257 | -80.4134
3.8261 | -80.4142
3.8264 | -80.415
3.8267 | -80.4153
3.827 | -80.4161
3.8273 | -80.4157
3.8277 | -80.4161
3.828 | -80.4161
3.8283 | -80.4159
3.8286 | -80.416
3.8289 | -80.4174
3.8293 | -80.4172
3.8296 | -80.4154
3.8299 | -80.4138
3.8302 | -80.4137
3.8306 | -80.4157
3.8309 | -80.4174
3.8312 | -80.4175
3.8315 | -80.417
3.8318 | -80.416
3.8322 | -80.4155
3.8325 | -80.4151
3.8328 | -80.4157
3.8331 | -80.416
3.8334 | -80.4161
3.8338 | -80.4158
3.8341 | -80.4151
3.8344 | -80.4151
3.8347 | -80.4157
3.8351 | -80.4156
3.8354 | -80.4144
3.8357 | -80.413
3.836 | -80.4132
3.8363 | -80.4165
3.8367 | -80.4189
3.837 | -80.4191
3.8373 | -80.4201
3.8376 | -80.4221
3.8379 | -80.423
3.8383 | -80.4235
3.8386 | -80.4242
3.8389 | -80.4259
3.8392 | -80.4263
3.8396 | -80.4266
3.8399 | -80.427
3.8402 | -80.4273
3.8405 | -80.4275
3.8408 | -80.4278
3.8412 | -80.4276
3.8415 | -80.4273
3.8418 | -80.4261
3.8421 | -80.4254
3.8424 | -80.4251
3.8428 | -80.4255
3.8431 | -80.4257
3.8434 | -80.426
3.8437 | -80.426
3.8441 | -80.426
3.8444 | -80.4256
3.8447 | -80.4252
3.845 | -80.426
3.8453 | -80.4256
3.8457 | -80.4247
3.846 | -80.4249
3.8463 | -80.4262
3.8466 | -80.4269
3.8469 | -80.4265
3.8473 | -80.4268
3.8476 | -80.4275
3.8479 | -80.4275
3.8482 | -80.4278
3.8486 | -80.4282
3.8489 | -80.4292
3.8492 | -80.4308
3.8495 | -80.4316
3.8498 | -80.4334
3.8502 | -80.4343
3.8505 | -80.4346
3.8508 | -80.4353
3.8511 | -80.4351
3.8514 | -80.4351
3.8518 | -80.4359
3.8521 | -80.4363
3.8524 | -80.4357
3.8527 | -80.4365
3.8531 | -80.4372
3.8534 | -80.4365
3.8537 | -80.4342
3.854 | -80.4331
3.8543 | -80.4329
3.8547 | -80.4337
3.855 | -80.4348
3.8553 | -80.4349
3.8556 | -80.4351
3.8559 | -80.4355
3.8563 | -80.437
3.8566 | -80.4365
3.8569 | -80.4384
3.8572 | -80.4394
3.8575 | -80.44
3.8579 | -80.44
3.8582 | -80.4397
3.8585 | -80.4407
3.8588 | -80.4418
3.8592 | -80.4423
3.8595 | -80.4432
3.8598 | -80.443
3.8601 | -80.4424
3.8604 | -80.4424
3.8608 | -80.4434
3.8611 | -80.4427
3.8614 | -80.4421
3.8617 | -80.4419
3.862 | -80.4402
3.8624 | -80.4391
3.8627 | -80.4389
3.863 | -80.4389
3.8633 | -80.4388
3.8637 | -80.4384
3.864 | -80.4382
3.8643 | -80.4377
3.8646 | -80.4371
3.8649 | -80.4369
3.8653 | -80.4363
3.8656 | -80.4351
3.8659 | -80.4357
3.8662 | -80.4359
3.8665 | -80.4353
3.8669 | -80.4356
3.8672 | -80.4375
3.8675 | -80.4386
3.8678 | -80.4386
3.8682 | -80.44
3.8685 | -80.4407
3.8688 | -80.44
3.8691 | -80.4393
3.8694 | -80.4388
3.8698 | -80.4362
3.8701 | -80.4343
3.8704 | -80.4351
3.8707 | -80.4334
3.871 | -80.4317
3.8714 | -80.4303
3.8717 | -80.4275
3.872 | -80.4283
3.8723 | -80.4308
3.8727 | -80.4306
3.873 | -80.431
3.8733 | -80.4315
3.8736 | -80.4315
3.8739 | -80.432
3.8743 | -80.4331
3.8746 | -80.4324
3.8749 | -80.4322
3.8752 | -80.4319
3.8755 | -80.4318
3.8759 | -80.4315
3.8762 | -80.4314
3.8765 | -80.4318
3.8768 | -80.4317
3.8771 | -80.4305
3.8775 | -80.429
3.8778 | -80.4296
3.8781 | -80.4299
3.8784 | -80.429
3.8788 | -80.4283
3.8791 | -80.4283
3.8794 | -80.4295
3.8797 | -80.4306
3.88 | -80.4302
3.8804 | -80.4297
3.8807 | -80.4296
3.881 | -80.4298
3.8813 | -80.43
3.8816 | -80.4306
3.882 | -80.4306
3.8823 | -80.4304
3.8826 | -80.4301
3.8829 | -80.4299
3.8833 | -80.4289
3.8836 | -80.428
3.8839 | -80.4274
3.8842 | -80.4277
3.8845 | -80.4266
3.8849 | -80.426
3.8852 | -80.4262
3.8855 | -80.4275
3.8858 | -80.4303
3.8861 | -80.4319
3.8865 | -80.431
3.8868 | -80.4317
3.8871 | -80.4313
3.8874 | -80.4308
3.8878 | -80.4312
3.8881 | -80.4339
3.8884 | -80.4358
3.8887 | -80.4364
3.889 | -80.4396
3.8894 | -80.44
3.8897 | -80.4395
3.89 | -80.4393
3.8903 | -80.4395
3.8906 | -80.4401
3.891 | -80.4412
3.8913 | -80.4416
3.8916 | -80.4394
3.8919 | -80.439
3.8923 | -80.4385
3.8926 | -80.4383
3.8929 | -80.4379
3.8932 | -80.4373
3.8935 | -80.4369
3.8939 | -80.436
3.8942 | -80.4355
3.8945 | -80.4357
3.8948 | -80.4345
3.8951 | -80.4321
3.8955 | -80.4311
3.8958 | -80.4301
3.8961 | -80.4298
3.8964 | -80.4294
3.8968 | -80.4291
3.8971 | -80.4294
3.8974 | -80.4292
3.8977 | -80.4295
3.898 | -80.429
3.8984 | -80.4283
3.8987 | -80.4265
3.899 | -80.4253
3.8993 | -80.4246
3.8996 | -80.4239
3.9 | -80.4227
3.9003 | -80.4241
3.9006 | -80.4238
3.9009 | -80.4238
3.9012 | -80.4231
3.9016 | -80.4214
3.9019 | -80.4197
3.9022 | -80.4189
3.9025 | -80.4186
3.9029 | -80.4182
3.9032 | -80.418
3.9035 | -80.4173
3.9038 | -80.4177
3.9041 | -80.4187
3.9045 | -80.4183
3.9048 | -80.4177
3.9051 | -80.4188
3.9054 | -80.4189
3.9057 | -80.4188
3.9061 | -80.4182
3.9064 | -80.4179
3.9067 | -80.4185
3.907 | -80.4186
3.9074 | -80.4184
3.9077 | -80.419
3.908 | -80.4197
3.9083 | -80.4194
3.9086 | -80.4187
3.909 | -80.4185
3.9093 | -80.4186
3.9096 | -80.4185
3.9099 | -80.4163
3.9102 | -80.4137
3.9106 | -80.4129
3.9109 | -80.4127
3.9112 | -80.4126
3.9115 | -80.4114
3.9119 | -80.4092
3.9122 | -80.4088
3.9125 | -80.4092
3.9128 | -80.41
3.9131 | -80.4111
3.9135 | -80.4111
3.9138 | -80.411
3.9141 | -80.4106
3.9144 | -80.4094
3.9147 | -80.4094
3.9151 | -80.4093
3.9154 | -80.4087
3.9157 | -80.4086
3.916 | -80.4081
3.9164 | -80.4065
3.9167 | -80.4049
3.917 | -80.4033
3.9173 | -80.4019
3.9176 | -80.4018
3.918 | -80.4014
3.9183 | -80.4005
3.9186 | -80.3988
3.9189 | -80.3981
3.9192 | -80.398
3.9196 | -80.3983
3.9199 | -80.3976
3.9202 | -80.3955
3.9205 | -80.3947
3.9209 | -80.3953
3.9212 | -80.3951
3.9215 | -80.3952
3.9218 | -80.3958
3.9221 | -80.3958
3.9225 | -80.3947
3.9228 | -80.3944
3.9231 | -80.3954
3.9234 | -80.395
3.9237 | -80.3943
3.9241 | -80.3938
3.9244 | -80.3936
3.9247 | -80.3925
3.925 | -80.3915
3.9253 | -80.3905
3.9257 | -80.3897
3.926 | -80.3885
3.9263 | -80.3864
3.9266 | -80.3856
3.927 | -80.3857
3.9273 | -80.3849
3.9276 | -80.3854
3.9279 | -80.3857
3.9282 | -80.3845
3.9286 | -80.3835
3.9289 | -80.3832
3.9292 | -80.3832
3.9295 | -80.383
3.9298 | -80.3825
3.9302 | -80.3817
3.9305 | -80.3808
3.9308 | -80.3807
3.9311 | -80.3801
3.9315 | -80.3805
3.9318 | -80.3809
3.9321 | -80.3806
3.9324 | -80.3807
3.9327 | -80.3794
3.9331 | -80.3801
3.9334 | -80.3821
3.9337 | -80.3816
3.934 | -80.3808
3.9343 | -80.3808
3.9347 | -80.3802
3.935 | -80.3796
3.9353 | -80.379
3.9356 | -80.3792
3.936 | -80.379
3.9363 | -80.3789
3.9366 | -80.3777
3.9369 | -80.3767
3.9372 | -80.3762
3.9376 | -80.3752
3.9379 | -80.3738
3.9382 | -80.3723
3.9385 | -80.3721
3.9388 | -80.3712
3.9392 | -80.3701
3.9395 | -80.369
3.9398 | -80.3685
3.9401 | -80.3676
3.9405 | -80.3671
3.9408 | -80.3681
3.9411 | -80.3676
3.9414 | -80.3668
3.9417 | -80.3661
3.9421 | -80.3652
3.9424 | -80.3638
3.9427 | -80.3632
3.943 | -80.3624
3.9433 | -80.3607
3.9437 | -80.3568
3.944 | -80.3542
3.9443 | -80.3518
3.9446 | -80.3487
3.9449 | -80.3478
3.9453 | -80.3479
3.9456 | -80.3469
3.9459 | -80.3473
3.9462 | -80.3467
3.9466 | -80.3461
3.9469 | -80.3456
3.9472 | -80.3446
3.9475 | -80.3441
3.9478 | -80.344
3.9482 | -80.3432
3.9485 | -80.3412
3.9488 | -80.3401
3.9491 | -80.3403
3.9494 | -80.3393
3.9498 | -80.3398
3.9501 | -80.3404
3.9504 | -80.3399
3.9507 | -80.3394
3.9511 | -80.3395
3.9514 | -80.339
3.9517 | -80.3395
3.952 | -80.3409
3.9523 | -80.3424
3.9527 | -80.3434
3.953 | -80.3423
3.9533 | -80.34
3.9536 | -80.3366
3.9539 | -80.3324
3.9543 | -80.3318
3.9546 | -80.3303
3.9549 | -80.3294
3.9552 | -80.3295
3.9556 | -80.3303
3.9559 | -80.33
3.9562 | -80.3288
3.9565 | -80.3281
3.9568 | -80.3276
3.9572 | -80.3264
3.9575 | -80.326
3.9578 | -80.325
3.9581 | -80.3238
3.9584 | -80.3237
3.9588 | -80.3235
3.9591 | -80.3225
3.9594 | -80.3217
3.9597 | -80.3206
3.9601 | -80.3191
3.9604 | -80.3188
3.9607 | -80.318
3.961 | -80.3163
3.9613 | -80.3144
3.9617 | -80.3137
3.962 | -80.3124
3.9623 | -80.3109
3.9626 | -80.3103
3.9629 | -80.3092
3.9633 | -80.3086
3.9636 | -80.3082
3.9639 | -80.3076
3.9642 | -80.3072
3.9646 | -80.3064
3.9649 | -80.3057
3.9652 | -80.3056
3.9655 | -80.305
3.9658 | -80.3033
3.9662 | -80.3026
3.9665 | -80.3029
3.9668 | -80.3029
3.9671 | -80.3021
3.9674 | -80.3018
3.9678 | -80.3021
3.9681 | -80.3017
3.9684 | -80.3011
3.9687 | -80.3001
3.9691 | -80.3005
3.9694 | -80.2988
3.9697 | -80.2967
3.97 | -80.2958
3.9703 | -80.2957
3.9707 | -80.2948
3.971 | -80.2943
3.9713 | -80.2959
3.9716 | -80.2981
3.9719 | -80.2975
3.9723 | -80.297
3.9726 | -80.2968
3.9729 | -80.2965
3.9732 | -80.296
3.9735 | -80.2961
3.9739 | -80.2967
3.9742 | -80.2966
3.9745 | -80.296
3.9748 | -80.2954
3.9752 | -80.2945
3.9755 | -80.293
3.9758 | -80.2929
3.9761 | -80.293
3.9764 | -80.2929
3.9768 | -80.2924
3.9771 | -80.2918
3.9774 | -80.2921
3.9777 | -80.2932
3.978 | -80.293
3.9784 | -80.2919
3.9787 | -80.2901
3.979 | -80.2891
3.9793 | -80.2881
3.9797 | -80.2873
3.98 | -80.2872
3.9803 | -80.2868
3.9806 | -80.2861
3.9809 | -80.2843
3.9813 | -80.2828
3.9816 | -80.2823
3.9819 | -80.2811
3.9822 | -80.2807
3.9825 | -80.2801
3.9829 | -80.2803
3.9832 | -80.2822
3.9835 | -80.2829
3.9838 | -80.2821
3.9842 | -80.2798
3.9845 | -80.278
3.9848 | -80.2772
3.9851 | -80.2767
3.9854 | -80.2764
3.9858 | -80.2761
3.9861 | -80.2767
3.9864 | -80.2752
3.9867 | -80.2761
3.987 | -80.2756
3.9874 | -80.2757
3.9877 | -80.2749
3.988 | -80.2739
3.9883 | -80.2741
3.9887 | -80.2725
3.989 | -80.2717
3.9893 | -80.2714
3.9896 | -80.271
3.9899 | -80.2709
3.9903 | -80.2711
3.9906 | -80.27
3.9909 | -80.2692
3.9912 | -80.2695
3.9915 | -80.2693
3.9919 | -80.269
3.9922 | -80.2687
3.9925 | -80.2681
3.9928 | -80.2679
3.9931 | -80.2677
3.9935 | -80.2673
3.9938 | -80.2672
3.9941 | -80.266
3.9944 | -80.266
3.9948 | -80.266
3.9951 | -80.2652
3.9954 | -80.265
3.9957 | -80.2643
3.996 | -80.2641
3.9964 | -80.2638
3.9967 | -80.2634
3.997 | -80.2626
3.9973 | -80.2618
3.9976 | -80.2618
3.998 | -80.2615
3.9983 | -80.2611
3.9986 | -80.2611
3.9989 | -80.2607
3.9993 | -80.2606
3.9996 | -80.2605
3.9999 | -80.2602
4.0002 | -80.2613
4.0005 | -80.2618
4.0009 | -80.2615
4.0012 | -80.2615
4.0015 | -80.2616
4.0018 | -80.2613
4.0021 | -80.2608
4.0025 | -80.2613
4.0028 | -80.2622
4.0031 | -80.2622
4.0034 | -80.2618
4.0038 | -80.2615
4.0041 | -80.2611
4.0044 | -80.2613
4.0047 | -80.2613
4.005 | -80.2615
4.0054 | -80.2614
4.0057 | -80.2608
4.006 | -80.2598
4.0063 | -80.2595
4.0066 | -80.2595
4.007 | -80.2596
4.0073 | -80.2613
4.0076 | -80.2626
4.0079 | -80.2625
4.0083 | -80.2632
4.0086 | -80.2631
4.0089 | -80.2634
4.0092 | -80.2642
4.0095 | -80.2656
4.0099 | -80.2656
4.0102 | -80.2653
4.0105 | -80.2647
4.0108 | -80.264
4.0111 | -80.264
4.0115 | -80.2641
4.0118 | -80.2644
4.0121 | -80.2636
4.0124 | -80.2631
4.0127 | -80.2636
4.0131 | -80.2647
4.0134 | -80.2641
4.0137 | -80.2649
4.014 | -80.2657
4.0144 | -80.2663
4.0147 | -80.2663
4.015 | -80.2663
4.0153 | -80.2662
4.0156 | -80.2659
4.016 | -80.2659
4.0163 | -80.2673
4.0166 | -80.2677
4.0169 | -80.2676
4.0172 | -80.2671
4.0176 | -80.2675
4.0179 | -80.2677
4.0182 | -80.2687
4.0185 | -80.2695
4.0189 | -80.2707
4.0192 | -80.2716
4.0195 | -80.27
4.0198 | -80.2686
4.0201 | -80.268
4.0205 | -80.2672
4.0208 | -80.2669
4.0211 | -80.2672
4.0214 | -80.267
4.0217 | -80.2656
4.0221 | -80.2641
4.0224 | -80.2639
4.0227 | -80.264
4.023 | -80.2632
4.0234 | -80.2633
4.0237 | -80.2639
4.024 | -80.2651
4.0243 | -80.2644
4.0246 | -80.2652
4.025 | -80.2664
4.0253 | -80.2677
4.0256 | -80.2694
4.0259 | -80.27
4.0262 | -80.2705
4.0266 | -80.271
4.0269 | -80.2702
4.0272 | -80.2685
4.0275 | -80.2683
4.0279 | -80.2688
4.0282 | -80.2694
4.0285 | -80.2698
4.0288 | -80.2701
4.0291 | -80.27
4.0295 | -80.2699
4.0298 | -80.2707
4.0301 | -80.2718
4.0304 | -80.2721
4.0307 | -80.2722
4.0311 | -80.2724
4.0314 | -80.2738
4.0317 | -80.2748
4.032 | -80.275
4.0324 | -80.2745
4.0327 | -80.2737
4.033 | -80.2718
4.0333 | -80.272
4.0336 | -80.2722
4.034 | -80.272
4.0343 | -80.2724
4.0346 | -80.2727
4.0349 | -80.273
4.0352 | -80.2735
4.0356 | -80.274
4.0359 | -80.2741
4.0362 | -80.2744
4.0365 | -80.2766
4.0369 | -80.2799
4.0372 | -80.28
4.0375 | -80.2803
4.0378 | -80.2802
4.0381 | -80.2804
4.0385 | -80.2805
4.0388 | -80.2798
4.0391 | -80.2796
4.0394 | -80.2802
4.0397 | -80.2802
4.0401 | -80.2806
4.0404 | -80.281
4.0407 | -80.2819
4.041 | -80.283
4.0413 | -80.2835
4.0417 | -80.2845
4.042 | -80.286
4.0423 | -80.2858
4.0426 | -80.2854
4.043 | -80.2877
4.0433 | -80.2878
4.0436 | -80.2874
4.0439 | -80.2871
4.0442 | -80.2871
4.0446 | -80.2875
4.0449 | -80.2873
4.0452 | -80.2873
4.0455 | -80.2886
4.0458 | -80.2889
4.0462 | -80.2902
4.0465 | -80.291
4.0468 | -80.2916
4.0471 | -80.2912
4.0475 | -80.2915
4.0478 | -80.2918
4.0481 | -80.2918
4.0484 | -80.2921
4.0487 | -80.2922
4.0491 | -80.2925
4.0494 | -80.2928
4.0497 | -80.2924
4.05 | -80.2923
4.0503 | -80.2922
4.0507 | -80.2931
4.051 | -80.2937
4.0513 | -80.2937
4.0516 | -80.2945
4.052 | -80.295
4.0523 | -80.2965
4.0526 | -80.2972
4.0529 | -80.2978
4.0532 | -80.298
4.0536 | -80.2979
4.0539 | -80.2991
4.0542 | -80.2996
4.0545 | -80.301
4.0548 | -80.3019
4.0552 | -80.3024
4.0555 | -80.3033
4.0558 | -80.3045
4.0561 | -80.3053
4.0565 | -80.3053
4.0568 | -80.3047
4.0571 | -80.3047
4.0574 | -80.3057
4.0577 | -80.3079
4.0581 | -80.3098
4.0584 | -80.3102
4.0587 | -80.3109
4.059 | -80.3112
4.0593 | -80.311
4.0597 | -80.3113
4.06 | -80.3114
4.0603 | -80.3126
4.0606 | -80.3144
4.061 | -80.3154
4.0613 | -80.316
4.0616 | -80.3167
4.0619 | -80.3172
4.0622 | -80.3179
4.0626 | -80.3196
4.0629 | -80.3213
4.0632 | -80.3222
4.0635 | -80.3244
4.0638 | -80.3253
4.0642 | -80.3268
4.0645 | -80.3276
4.0648 | -80.3285
4.0651 | -80.3304
4.0654 | -80.3307
4.0658 | -80.3314
4.0661 | -80.332
4.0664 | -80.3328
4.0667 | -80.3336
4.0671 | -80.3343
4.0674 | -80.335
4.0677 | -80.3358
4.068 | -80.3365
4.0683 | -80.3371
4.0687 | -80.3378
4.069 | -80.3383
4.0693 | -80.3393
4.0696 | -80.3399
4.0699 | -80.3419
4.0703 | -80.3428
4.0706 | -80.3435
4.0709 | -80.3441
4.0712 | -80.3447
4.0716 | -80.3451
4.0719 | -80.346
4.0722 | -80.3466
4.0725 | -80.348
4.0728 | -80.3488
4.0732 | -80.3507
4.0735 | -80.3519
4.0738 | -80.3529
4.0741 | -80.3537
4.0744 | -80.3549
4.0748 | -80.3561
4.0751 | -80.3571
4.0754 | -80.3585
4.0757 | -80.3595
4.0761 | -80.3607
4.0764 | -80.3619
4.0767 | -80.3626
4.077 | -80.3637
4.0773 | -80.3655
4.0777 | -80.3676
4.078 | -80.3691
4.0783 | -80.3702
4.0786 | -80.3717
4.0789 | -80.374
4.0793 | -80.3755
4.0796 | -80.3771
4.0799 | -80.3783
4.0802 | -80.3796
4.0806 | -80.3806
4.0809 | -80.382
4.0812 | -80.3831
4.0815 | -80.3841
4.0818 | -80.3846
4.0822 | -80.3851
4.0825 | -80.3861
4.0828 | -80.3892
4.0831 | -80.3912
4.0834 | -80.3923
4.0838 | -80.394
4.0841 | -80.3953
4.0844 | -80.3961
4.0847 | -80.3976
4.085 | -80.3986
4.0854 | -80.3986
4.0857 | -80.3984
4.086 | -80.399
4.0863 | -80.3997
4.0867 | -80.4006
4.087 | -80.4018
4.0873 | -80.4032
4.0876 | -80.4044
4.0879 | -80.4071
4.0883 | -80.4095
4.0886 | -80.4121
4.0889 | -80.4139
4.0892 | -80.4152
4.0895 | -80.4183
4.0899 | -80.4198
4.0902 | -80.421
4.0905 | -80.4224
4.0908 | -80.4234
4.0912 | -80.4233
4.0915 | -80.4238
4.0918 | -80.4252
4.0921 | -80.4264
4.0924 | -80.4271
4.0928 | -80.4274
4.0931 | -80.4278
4.0934 | -80.4281
4.0937 | -80.4298
4.094 | -80.4314
4.0944 | -80.4334
4.0947 | -80.4354
4.095 | -80.4368
4.0953 | -80.4381
4.0957 | -80.4407
4.096 | -80.4426
4.0963 | -80.4431
4.0966 | -80.4437
4.0969 | -80.4445
4.0973 | -80.4453
4.0976 | -80.4432
4.0979 | -80.443
4.0982 | -80.4439
4.0985 | -80.4452
4.0989 | -80.4473
4.0992 | -80.45
4.0995 | -80.4517
4.0998 | -80.453
4.1002 | -80.4551
4.1005 | -80.4569
4.1008 | -80.4584
4.1011 | -80.4589
4.1014 | -80.4596
4.1018 | -80.4609
4.1021 | -80.4618
4.1024 | -80.4635
4.1027 | -80.4657
4.103 | -80.4662
4.1034 | -80.4661
4.1037 | -80.4674
4.104 | -80.469
4.1043 | -80.4702
4.1047 | -80.4715
4.105 | -80.4741
4.1053 | -80.4777
4.1056 | -80.481
4.1059 | -80.4831
4.1063 | -80.4847
4.1066 | -80.4835
4.1069 | -80.4845
4.1072 | -80.4852
4.1075 | -80.4874
4.1079 | -80.4895
4.1082 | -80.4901
4.1085 | -80.491
4.1088 | -80.4927
4.1091 | -80.4943
4.1095 | -80.4958
4.1098 | -80.4975
4.1101 | -80.4987
4.1104 | -80.4999
4.1108 | -80.5011
4.1111 | -80.5017
4.1114 | -80.5044
4.1117 | -80.507
4.112 | -80.5076
4.1124 | -80.5091
4.1127 | -80.5112
4.113 | -80.5126
4.1133 | -80.514
4.1136 | -80.5151
4.114 | -80.5165
4.1143 | -80.5173
4.1146 | -80.5179
4.1149 | -80.5166
4.1153 | -80.5157
4.1156 | -80.516
4.1159 | -80.5162
4.1162 | -80.5185
4.1165 | -80.5192
4.1169 | -80.5203
4.1172 | -80.5214
4.1175 | -80.5206
4.1178 | -80.5201
4.1181 | -80.5215
4.1185 | -80.5227
4.1188 | -80.5224
4.1191 | -80.5232
4.1194 | -80.5232
4.1198 | -80.5243
4.1201 | -80.5257
4.1204 | -80.5271
4.1207 | -80.5277
4.121 | -80.5281
4.1214 | -80.5298
4.1217 | -80.5311
4.122 | -80.5331
4.1223 | -80.5364
4.1226 | -80.5384
4.123 | -80.539
4.1233 | -80.5396
4.1236 | -80.5418
4.1239 | -80.5437
4.1243 | -80.545
4.1246 | -80.5457
4.1249 | -80.546
4.1252 | -80.5471
4.1255 | -80.5491
4.1259 | -80.5502
4.1262 | -80.5515
4.1265 | -80.5527
4.1268 | -80.5527
4.1271 | -80.5525
4.1275 | -80.5533
4.1278 | -80.5546
4.1281 | -80.5555
4.1284 | -80.5574
4.1288 | -80.5591
4.1291 | -80.5608
4.1294 | -80.5628
4.1297 | -80.5646
4.13 | -80.5667
4.1304 | -80.5688
4.1307 | -80.5713
4.131 | -80.5753
4.1313 | -80.5781
4.1316 | -80.5803
4.132 | -80.582
4.1323 | -80.5829
4.1326 | -80.5839
4.1329 | -80.5849
4.1332 | -80.5861
4.1336 | -80.5875
4.1339 | -80.5897
4.1342 | -80.5913
4.1345 | -80.5929
4.1349 | -80.5947
4.1352 | -80.5955
4.1355 | -80.5943
4.1358 | -80.5918
4.1361 | -80.5928
4.1365 | -80.594
4.1368 | -80.5952
4.1371 | -80.5963
4.1374 | -80.5977
4.1377 | -80.5991
4.1381 | -80.6003
4.1384 | -80.6002
4.1387 | -80.6005
4.139 | -80.6019
4.1394 | -80.6036
4.1397 | -80.6053
4.14 | -80.6077
4.1403 | -80.6094
4.1406 | -80.6106
4.141 | -80.6125
4.1413 | -80.6127
4.1416 | -80.6116
4.1419 | -80.6119
4.1422 | -80.6137
4.1426 | -80.6157
4.1429 | -80.6171
4.1432 | -80.6185
4.1435 | -80.6204
4.1439 | -80.622
4.1442 | -80.6236
4.1445 | -80.6246
4.1448 | -80.6259
4.1451 | -80.6278
4.1455 | -80.6291
4.1458 | -80.6301
4.1461 | -80.6312
4.1464 | -80.6322
4.1467 | -80.6328
4.1471 | -80.6338
4.1474 | -80.6355
4.1477 | -80.6371
4.148 | -80.6385
4.1484 | -80.6411
4.1487 | -80.6426
4.149 | -80.6427
4.1493 | -80.6458
4.1496 | -80.6492
4.15 | -80.6515
4.1503 | -80.6548
4.1506 | -80.657
4.1509 | -80.6579
4.1512 | -80.6583
4.1516 | -80.6586
4.1519 | -80.6595
4.1522 | -80.6603
4.1525 | -80.661
4.1528 | -80.6618
4.1532 | -80.6612
4.1535 | -80.6626
4.1538 | -80.6619
4.1541 | -80.6618
4.1545 | -80.6629
4.1548 | -80.6637
4.1551 | -80.6646
4.1554 | -80.6656
4.1557 | -80.6658
4.1561 | -80.667
4.1564 | -80.6696
4.1567 | -80.6723
4.157 | -80.6731
4.1573 | -80.6745
4.1577 | -80.6758
4.158 | -80.677
4.1583 | -80.6779
4.1586 | -80.6789
4.159 | -80.6803
4.1593 | -80.6819
4.1596 | -80.6826
4.1599 | -80.6834
4.1602 | -80.6839
4.1606 | -80.6842
4.1609 | -80.684
4.1612 | -80.6847
4.1615 | -80.6857
4.1618 | -80.6864
4.1622 | -80.6874
4.1625 | -80.6876
4.1628 | -80.688
4.1631 | -80.6897
4.1635 | -80.6907
4.1638 | -80.6913
4.1641 | -80.6922
4.1644 | -80.6945
4.1647 | -80.6963
4.1651 | -80.6963
4.1654 | -80.6968
4.1657 | -80.6971
4.166 | -80.6982
4.1663 | -80.6996
4.1667 | -80.7005
4.167 | -80.7016
4.1673 | -80.7028
4.1676 | -80.7031
4.168 | -80.7038
4.1683 | -80.7038
4.1686 | -80.7051
4.1689 | -80.7059
4.1692 | -80.7065
4.1696 | -80.7068
4.1699 | -80.7071
4.1702 | -80.7077
4.1705 | -80.7088
4.1708 | -80.7095
4.1712 | -80.7106
4.1715 | -80.7131
4.1718 | -80.714
4.1721 | -80.7147
4.1725 | -80.7156
4.1728 | -80.7147
4.1731 | -80.714
4.1734 | -80.7146
4.1737 | -80.7152
4.1741 | -80.7155
4.1744 | -80.7169
4.1747 | -80.718
4.175 | -80.7187
4.1753 | -80.7202
4.1757 | -80.7236
4.176 | -80.7253
4.1763 | -80.7267
4.1766 | -80.7273
4.1769 | -80.726
4.1773 | -80.7247
4.1776 | -80.7252
4.1779 | -80.7256
4.1782 | -80.7244
4.1786 | -80.7236
4.1789 | -80.7246
4.1792 | -80.7266
4.1795 | -80.7276
4.1798 | -80.7276
4.1802 | -80.7286
4.1805 | -80.7298
4.1808 | -80.731
4.1811 | -80.7317
4.1814 | -80.7325
4.1818 | -80.7331
4.1821 | -80.734
4.1824 | -80.7346
4.1827 | -80.7362
4.1831 | -80.7377
4.1834 | -80.7381
4.1837 | -80.7386
4.184 | -80.7395
4.1843 | -80.7394
4.1847 | -80.7387
4.185 | -80.7394
4.1853 | -80.739
4.1856 | -80.7389
4.1859 | -80.7392
4.1863 | -80.7389
4.1866 | -80.7383
4.1869 | -80.7387
4.1872 | -80.7394
4.1876 | -80.7404
4.1879 | -80.7408
4.1882 | -80.7416
4.1885 | -80.742
4.1888 | -80.7419
4.1892 | -80.742
4.1895 | -80.7424
4.1898 | -80.743
4.1901 | -80.7433
4.1904 | -80.744
4.1908 | -80.7444
4.1911 | -80.7451
4.1914 | -80.7474
4.1917 | -80.7499
4.1921 | -80.7507
4.1924 | -80.751
4.1927 | -80.7512
4.193 | -80.753
4.1933 | -80.7552
4.1937 | -80.7553
4.194 | -80.7549
4.1943 | -80.7529
4.1946 | -80.7511
4.1949 | -80.7496
4.1953 | -80.7484
4.1956 | -80.7481
4.1959 | -80.7481
4.1962 | -80.749
4.1966 | -80.7494
4.1969 | -80.7498
4.1972 | -80.7502
4.1975 | -80.7503
4.1978 | -80.7506
4.1982 | -80.7521
4.1985 | -80.7535
4.1988 | -80.7547
4.1991 | -80.7556
4.1994 | -80.7569
4.1998 | -80.756
4.2001 | -80.7561
4.2004 | -80.7566
4.2007 | -80.7581
4.201 | -80.7577
4.2014 | -80.7579
4.2017 | -80.7582
4.202 | -80.7587
4.2023 | -80.7597
4.2027 | -80.7602
4.203 | -80.7596
4.2033 | -80.7596
4.2036 | -80.7599
4.2039 | -80.7602
4.2043 | -80.7601
4.2046 | -80.7602
4.2049 | -80.762
4.2052 | -80.7628
4.2055 | -80.7623
4.2059 | -80.7624
4.2062 | -80.7627
4.2065 | -80.7625
4.2068 | -80.7621
4.2072 | -80.7629
4.2075 | -80.7639
4.2078 | -80.7638
4.2081 | -80.7645
4.2084 | -80.7657
4.2088 | -80.7676
4.2091 | -80.7686
4.2094 | -80.77
4.2097 | -80.7698
4.21 | -80.7701
4.2104 | -80.7704
4.2107 | -80.7696
4.211 | -80.7704
4.2113 | -80.7702
4.2117 | -80.7701
4.212 | -80.7702
4.2123 | -80.7701
4.2126 | -80.7704
4.2129 | -80.77
4.2133 | -80.7703
4.2136 | -80.7693
4.2139 | -80.7692
4.2142 | -80.769
4.2145 | -80.7674
4.2149 | -80.7658
4.2152 | -80.7636
4.2155 | -80.7631
4.2158 | -80.7642
4.2162 | -80.7645
4.2165 | -80.7643
4.2168 | -80.7645
4.2171 | -80.7635
4.2174 | -80.7633
4.2178 | -80.7635
4.2181 | -80.7635
4.2184 | -80.762
4.2187 | -80.7621
4.219 | -80.7617
4.2194 | -80.7632
4.2197 | -80.7637
4.22 | -80.7637
4.2203 | -80.7642
4.2207 | -80.7642
4.221 | -80.7648
4.2213 | -80.7651
4.2216 | -80.766
4.2219 | -80.7666
4.2223 | -80.7669
4.2226 | -80.7665
4.2229 | -80.7666
4.2232 | -80.7666
4.2235 | -80.7672
4.2239 | -80.767
4.2242 | -80.7672
4.2245 | -80.7669
4.2248 | -80.7675
4.2251 | -80.7683
4.2255 | -80.7678
4.2258 | -80.7676
4.2261 | -80.7685
4.2264 | -80.7685
4.2268 | -80.769
4.2271 | -80.7701
4.2274 | -80.7709
4.2277 | -80.7745
4.228 | -80.7768
4.2284 | -80.7781
4.2287 | -80.7798
4.229 | -80.7803
4.2293 | -80.7802
4.2296 | -80.7797
4.23 | -80.7794
4.2303 | -80.779
4.2306 | -80.7789
4.2309 | -80.779
4.2313 | -80.7793
4.2316 | -80.7793
4.2319 | -80.7789
4.2322 | -80.7792
4.2325 | -80.7805
4.2329 | -80.7813
4.2332 | -80.7819
4.2335 | -80.7828
4.2338 | -80.7831
4.2341 | -80.782
4.2345 | -80.7816
4.2348 | -80.7815
4.2351 | -80.7815
4.2354 | -80.7813
4.2358 | -80.78
4.2361 | -80.7806
4.2364 | -80.7814
4.2367 | -80.7825
4.237 | -80.7827
4.2374 | -80.7826
4.2377 | -80.7832
4.238 | -80.7852
4.2383 | -80.7885
4.2386 | -80.7921
4.239 | -80.7915
4.2393 | -80.7904
4.2396 | -80.7901
4.2399 | -80.7897
4.2403 | -80.7888
4.2406 | -80.7879
4.2409 | -80.7875
4.2412 | -80.7874
4.2415 | -80.7876
4.2419 | -80.7882
4.2422 | -80.7881
4.2425 | -80.7888
4.2428 | -80.7899
4.2431 | -80.7897
4.2435 | -80.7896
4.2438 | -80.7887
4.2441 | -80.788
4.2444 | -80.7862
4.2447 | -80.7861
4.2451 | -80.7868
4.2454 | -80.7865
4.2457 | -80.7872
4.246 | -80.7884
4.2464 | -80.7893
4.2467 | -80.7908
4.247 | -80.7909
4.2473 | -80.7909
4.2476 | -80.7918
4.248 | -80.7916
4.2483 | -80.7922
4.2486 | -80.7924
4.2489 | -80.7923
4.2492 | -80.7921
4.2496 | -80.791
4.2499 | -80.7904
4.2502 | -80.7901
4.2505 | -80.7896
4.2509 | -80.7902
4.2512 | -80.7901
4.2515 | -80.7884
4.2518 | -80.7877
4.2521 | -80.788
4.2525 | -80.7877
4.2528 | -80.7862
4.2531 | -80.7864
4.2534 | -80.7863
4.2537 | -80.7863
4.2541 | -80.7871
4.2544 | -80.788
4.2547 | -80.7887
4.255 | -80.7906
4.2554 | -80.792
4.2557 | -80.794
4.256 | -80.7943
4.2563 | -80.7947
4.2566 | -80.7955
4.257 | -80.7963
4.2573 | -80.7965
4.2576 | -80.7967
4.2579 | -80.7973
4.2582 | -80.7971
4.2586 | -80.7969
4.2589 | -80.7969
4.2592 | -80.7973
4.2595 | -80.7981
4.2599 | -80.7987
4.2602 | -80.7988
4.2605 | -80.7993
4.2608 | -80.7997
4.2611 | -80.8002
4.2615 | -80.8018
4.2618 | -80.8021
4.2621 | -80.8011
4.2624 | -80.8008
4.2627 | -80.8006
4.2631 | -80.8014
4.2634 | -80.8022
4.2637 | -80.8026
4.264 | -80.8038
4.2644 | -80.8035
4.2647 | -80.8051
4.265 | -80.8064
4.2653 | -80.8094
4.2656 | -80.8097
4.266 | -80.8087
4.2663 | -80.8089
4.2666 | -80.8093
4.2669 | -80.8086
4.2672 | -80.8078
4.2676 | -80.8101
4.2679 | -80.8123
4.2682 | -80.8129
4.2685 | -80.8135
4.2688 | -80.8133
4.2692 | -80.8139
4.2695 | -80.8145
4.2698 | -80.815
4.2701 | -80.816
4.2705 | -80.8172
4.2708 | -80.8195
4.2711 | -80.8212
4.2714 | -80.8222
4.2717 | -80.8232
4.2721 | -80.8236
4.2724 | -80.8248
4.2727 | -80.8267
4.273 | -80.8282
4.2733 | -80.8302
4.2737 | -80.8319
4.274 | -80.8328
4.2743 | -80.8328
4.2746 | -80.8335
4.275 | -80.8355
4.2753 | -80.8372
4.2756 | -80.8379
4.2759 | -80.839
4.2762 | -80.8399
4.2766 | -80.8408
4.2769 | -80.842
4.2772 | -80.8413
4.2775 | -80.8416
4.2778 | -80.8429
4.2782 | -80.8441
4.2785 | -80.8447
4.2788 | -80.8456
4.2791 | -80.846
4.2795 | -80.8455
4.2798 | -80.846
4.2801 | -80.847
4.2804 | -80.8476
4.2807 | -80.8482
4.2811 | -80.8492
4.2814 | -80.8504
4.2817 | -80.8512
4.282 | -80.8521
4.2823 | -80.8536
4.2827 | -80.8552
4.283 | -80.8567
4.2833 | -80.8582
4.2836 | -80.8601
4.284 | -80.861
4.2843 | -80.8617
4.2846 | -80.8621
4.2849 | -80.8632
4.2852 | -80.8651
4.2856 | -80.8653
4.2859 | -80.8651
4.2862 | -80.8663
4.2865 | -80.8672
4.2868 | -80.8682
4.2872 | -80.8694
4.2875 | -80.8707
4.2878 | -80.8719
4.2881 | -80.873
4.2885 | -80.874
4.2888 | -80.875
4.2891 | -80.8756
4.2894 | -80.8767
4.2897 | -80.8776
4.2901 | -80.8788
4.2904 | -80.8797
4.2907 | -80.881
4.291 | -80.8815
4.2913 | -80.8819
4.2917 | -80.883
4.292 | -80.8846
4.2923 | -80.8855
4.2926 | -80.8848
4.2929 | -80.8842
4.2933 | -80.8864
4.2936 | -80.8871
4.2939 | -80.8878
4.2942 | -80.8896
4.2946 | -80.8913
4.2949 | -80.8928
4.2952 | -80.8932
4.2955 | -80.894
4.2958 | -80.8954
4.2962 | -80.8964
4.2965 | -80.8975
4.2968 | -80.8989
4.2971 | -80.9002
4.2974 | -80.9017
4.2978 | -80.9027
4.2981 | -80.904
4.2984 | -80.9048
4.2987 | -80.9049
4.2991 | -80.9066
4.2994 | -80.9087
4.2997 | -80.91
4.3 | -80.9111
4.3003 | -80.9116
4.3007 | -80.914
4.301 | -80.9158
4.3013 | -80.9166
4.3016 | -80.9182
4.3019 | -80.9193
4.3023 | -80.921
4.3026 | -80.9224
4.3029 | -80.9243
4.3032 | -80.926
4.3036 | -80.9279
4.3039 | -80.9293
4.3042 | -80.93
4.3045 | -80.9305
4.3048 | -80.932
4.3052 | -80.9326
4.3055 | -80.9332
4.3058 | -80.935
4.3061 | -80.9367
4.3064 | -80.9388
4.3068 | -80.9395
4.3071 | -80.9407
4.3074 | -80.9419
4.3077 | -80.9423
4.3081 | -80.9432
4.3084 | -80.945
4.3087 | -80.9482
4.309 | -80.951
4.3093 | -80.9529
4.3097 | -80.9547
4.31 | -80.9553
4.3103 | -80.9548
4.3106 | -80.9551
4.3109 | -80.9567
4.3113 | -80.9588
4.3116 | -80.962
4.3119 | -80.9637
4.3122 | -80.9649
4.3125 | -80.9673
4.3129 | -80.9703
4.3132 | -80.9716
4.3135 | -80.9724
4.3138 | -80.9735
4.3142 | -80.9758
4.3145 | -80.977
4.3148 | -80.9771
4.3151 | -80.9781
4.3154 | -80.9794
4.3158 | -80.9811
4.3161 | -80.9827
4.3164 | -80.9839
4.3167 | -80.9851
4.317 | -80.9876
4.3174 | -80.9892
4.3177 | -80.9903
4.318 | -80.992
4.3183 | -80.9937
4.3187 | -80.9949
4.319 | -80.9967
4.3193 | -80.9989
4.3196 | -81.0011
4.3199 | -81.0034
4.3203 | -81.0056
4.3206 | -81.0074
4.3209 | -81.0089
4.3212 | -81.0112
4.3215 | -81.0121
4.3219 | -81.0126
4.3222 | -81.0134
4.3225 | -81.0148
4.3228 | -81.0166
4.3232 | -81.0194
4.3235 | -81.0234
4.3238 | -81.0264
4.3241 | -81.0283
4.3244 | -81.0298
4.3248 | -81.0307
4.3251 | -81.0316
4.3254 | -81.0325
4.3257 | -81.0332
4.326 | -81.0342
4.3264 | -81.0354
4.3267 | -81.0368
4.327 | -81.0374
4.3273 | -81.0382
4.3277 | -81.04
4.328 | -81.0416
4.3283 | -81.0434
4.3286 | -81.0446
4.3289 | -81.0458
4.3293 | -81.0465
4.3296 | -81.045
4.3299 | -81.0468
4.3302 | -81.0477
4.3305 | -81.0493
4.3309 | -81.0509
4.3312 | -81.052
4.3315 | -81.053
4.3318 | -81.0542
4.3322 | -81.0565
4.3325 | -81.0575
4.3328 | -81.0584
4.3331 | -81.06
4.3334 | -81.0616
4.3338 | -81.064
4.3341 | -81.0655
4.3344 | -81.0668
4.3347 | -81.0689
4.335 | -81.0706
4.3354 | -81.0723
4.3357 | -81.0734
4.336 | -81.0749
4.3363 | -81.0773
4.3366 | -81.0788
4.337 | -81.0814
4.3373 | -81.0829
4.3376 | -81.0842
4.3379 | -81.0855
4.3383 | -81.088
4.3386 | -81.0886
4.3389 | -81.0898
4.3392 | -81.0909
4.3395 | -81.0919
4.3399 | -81.0936
4.3402 | -81.0956
4.3405 | -81.0968
4.3408 | -81.0977
4.3411 | -81.0997
4.3415 | -81.1014
4.3418 | -81.1028
4.3421 | -81.1038
4.3424 | -81.1051
4.3428 | -81.1077
4.3431 | -81.1092
4.3434 | -81.1098
4.3437 | -81.1106
4.344 | -81.1113
4.3444 | -81.1122
4.3447 | -81.1135
4.345 | -81.1151
4.3453 | -81.117
4.3456 | -81.118
4.346 | -81.1198
4.3463 | -81.1214
4.3466 | -81.1229
4.3469 | -81.124
4.3473 | -81.1244
4.3476 | -81.1256
4.3479 | -81.1274
4.3482 | -81.1289
4.3485 | -81.1307
4.3489 | -81.1336
4.3492 | -81.1353
4.3495 | -81.1363
4.3498 | -81.1375
4.3501 | -81.1375
4.3505 | -81.1383
4.3508 | -81.1396
4.3511 | -81.141
4.3514 | -81.1427
4.3518 | -81.1457
4.3521 | -81.1485
4.3524 | -81.1498
4.3527 | -81.1503
4.353 | -81.1508
4.3534 | -81.1517
4.3537 | -81.1531
4.354 | -81.1544
4.3543 | -81.1558
4.3546 | -81.157
4.355 | -81.158
4.3553 | -81.1592
4.3556 | -81.161
4.3559 | -81.1622
4.3563 | -81.1638
4.3566 | -81.1649
4.3569 | -81.1652
4.3572 | -81.1651
4.3575 | -81.1655
4.3579 | -81.1668
4.3582 | -81.1683
4.3585 | -81.1689
4.3588 | -81.1699
4.3591 | -81.171
4.3595 | -81.1719
4.3598 | -81.173
4.3601 | -81.1728
4.3604 | -81.1717
4.3607 | -81.1727
4.3611 | -81.1741
4.3614 | -81.1751
4.3617 | -81.1767
4.362 | -81.1782
4.3624 | -81.1796
4.3627 | -81.1806
4.363 | -81.1811
4.3633 | -81.1817
4.3636 | -81.1825
4.364 | -81.1831
4.3643 | -81.1841
4.3646 | -81.1856
4.3649 | -81.1852
4.3652 | -81.1837
4.3656 | -81.1828
4.3659 | -81.1836
4.3662 | -81.1845
4.3665 | -81.1857
4.3669 | -81.1862
4.3672 | -81.1868
4.3675 | -81.1872
4.3678 | -81.1878
4.3681 | -81.1881
4.3685 | -81.1882
4.3688 | -81.1889
4.3691 | -81.1893
4.3694 | -81.1886
4.3697 | -81.1887
4.3701 | -81.1895
4.3704 | -81.1894
4.3707 | -81.1889
4.371 | -81.1888
4.3714 | -81.1899
4.3717 | -81.1907
4.372 | -81.1916
4.3723 | -81.1928
4.3726 | -81.1944
4.373 | -81.1954
4.3733 | -81.1963
4.3736 | -81.1974
4.3739 | -81.1982
4.3742 | -81.1991
4.3746 | -81.1996
4.3749 | -81.2006
4.3752 | -81.2017
4.3755 | -81.2037
4.3759 | -81.2018
4.3762 | -81.2028
4.3765 | -81.2028
4.3768 | -81.2037
4.3771 | -81.2049
4.3775 | -81.2064
4.3778 | -81.2073
4.3781 | -81.208
4.3784 | -81.2094
4.3787 | -81.2105
4.3791 | -81.2115
4.3794 | -81.213
4.3797 | -81.2136
4.38 | -81.2143
4.3803 | -81.2153
4.3807 | -81.2161
4.381 | -81.2167
4.3813 | -81.2171
4.3816 | -81.2182
4.382 | -81.2189
4.3823 | -81.2189
4.3826 | -81.2179
4.3829 | -81.2181
4.3832 | -81.2184
4.3836 | -81.2184
4.3839 | -81.2191
4.3842 | -81.2198
4.3845 | -81.2203
4.3848 | -81.2202
4.3852 | -81.2199
4.3855 | -81.2199
4.3858 | -81.2212
4.3861 | -81.2228
4.3865 | -81.2234
4.3868 | -81.2241
4.3871 | -81.2253
4.3874 | -81.2261
4.3877 | -81.2272
4.3881 | -81.229
4.3884 | -81.2294
4.3887 | -81.2305
4.389 | -81.2305
4.3893 | -81.2309
4.3897 | -81.2315
4.39 | -81.232
4.3903 | -81.2314
4.3906 | -81.2318
4.391 | -81.233
4.3913 | -81.2341
4.3916 | -81.2348
4.3919 | -81.2368
4.3922 | -81.2389
4.3926 | -81.2415
4.3929 | -81.2434
4.3932 | -81.244
4.3935 | -81.2448
4.3938 | -81.2445
4.3942 | -81.2446
4.3945 | -81.2446
4.3948 | -81.2445
4.3951 | -81.2448
4.3955 | -81.2453
4.3958 | -81.2455
4.3961 | -81.2444
4.3964 | -81.2447
4.3967 | -81.2459
4.3971 | -81.246
4.3974 | -81.2469
4.3977 | -81.2467
4.398 | -81.2461
4.3983 | -81.2473
4.3987 | -81.2485
4.399 | -81.247
4.3993 | -81.2463
4.3996 | -81.2466
4.4 | -81.246
4.4003 | -81.245
4.4006 | -81.2453
4.4009 | -81.2451
4.4012 | -81.2447
4.4016 | -81.2448
4.4019 | -81.2461
4.4022 | -81.2471
4.4025 | -81.2476
4.4028 | -81.2483
4.4032 | -81.2482
4.4035 | -81.2472
4.4038 | -81.248
4.4041 | -81.2497
4.4044 | -81.2507
4.4048 | -81.2507
4.4051 | -81.2506
4.4054 | -81.25
4.4057 | -81.2492
4.4061 | -81.2499
4.4064 | -81.2503
4.4067 | -81.2502
4.407 | -81.2511
4.4073 | -81.2491
4.4077 | -81.2475
4.408 | -81.2483
4.4083 | -81.2488
4.4086 | -81.2486
4.4089 | -81.2486
4.4093 | -81.2476
4.4096 | -81.2463
4.4099 | -81.2461
4.4102 | -81.2464
4.4106 | -81.2471
4.4109 | -81.25
4.4112 | -81.2524
4.4115 | -81.2528
4.4118 | -81.2526
4.4122 | -81.2521
4.4125 | -81.2515
4.4128 | -81.2514
4.4131 | -81.2509
4.4134 | -81.2507
4.4138 | -81.2494
4.4141 | -81.2479
4.4144 | -81.2471
4.4147 | -81.2466
4.4151 | -81.2474
4.4154 | -81.2493
4.4157 | -81.2498
4.416 | -81.2507
4.4163 | -81.251
4.4167 | -81.2505
4.417 | -81.2502
4.4173 | -81.251
4.4176 | -81.2517
4.4179 | -81.2518
4.4183 | -81.2513
4.4186 | -81.2511
4.4189 | -81.2505
4.4192 | -81.25
4.4196 | -81.2498
4.4199 | -81.2497
4.4202 | -81.2494
4.4205 | -81.247
4.4208 | -81.2458
4.4212 | -81.2458
4.4215 | -81.2464
4.4218 | -81.246
4.4221 | -81.2446
4.4224 | -81.2429
4.4228 | -81.2421
4.4231 | -81.2408
4.4234 | -81.2401
4.4237 | -81.2398
4.4241 | -81.2401
4.4244 | -81.241
4.4247 | -81.2396
4.425 | -81.2385
4.4253 | -81.238
4.4257 | -81.2379
4.426 | -81.2375
4.4263 | -81.2372
4.4266 | -81.2368
4.4269 | -81.2366
4.4273 | -81.236
4.4276 | -81.2355
4.4279 | -81.2349
4.4282 | -81.2342
4.4285 | -81.2337
4.4289 | -81.2333
4.4292 | -81.2323
4.4295 | -81.2308
4.4298 | -81.2302
4.4302 | -81.2299
4.4305 | -81.2295
4.4308 | -81.2286
4.4311 | -81.2276
4.4314 | -81.2272
4.4318 | -81.2259
4.4321 | -81.2239
4.4324 | -81.2232
4.4327 | -81.2222
4.433 | -81.2223
4.4334 | -81.2243
4.4337 | -81.2277
4.434 | -81.2263
4.4343 | -81.2241
4.4347 | -81.2229
4.435 | -81.2217
4.4353 | -81.2207
4.4356 | -81.2208
4.4359 | -81.2205
4.4363 | -81.2213
4.4366 | -81.2212
4.4369 | -81.2194
4.4372 | -81.218
4.4375 | -81.2173
4.4379 | -81.2175
4.4382 | -81.2177
4.4385 | -81.2172
4.4388 | -81.2163
4.4392 | -81.2161
4.4395 | -81.2184
4.4398 | -81.2189
4.4401 | -81.2182
4.4404 | -81.2183
4.4408 | -81.218
4.4411 | -81.2169
4.4414 | -81.2168
4.4417 | -81.2163
4.442 | -81.2153
4.4424 | -81.2147
4.4427 | -81.2141
4.443 | -81.2135
4.4433 | -81.2129
4.4437 | -81.2123
4.444 | -81.2116
4.4443 | -81.2112
4.4446 | -81.2105
4.4449 | -81.21
4.4453 | -81.21
4.4456 | -81.2095
4.4459 | -81.2078
4.4462 | -81.2068
4.4465 | -81.2065
4.4469 | -81.2068
4.4472 | -81.208
4.4475 | -81.2079
4.4478 | -81.2069
4.4482 | -81.2061
4.4485 | -81.2059
4.4488 | -81.2054
4.4491 | -81.2049
4.4494 | -81.2045
4.4498 | -81.2038
4.4501 | -81.2024
4.4504 | -81.2011
4.4507 | -81.1995
4.451 | -81.1989
4.4514 | -81.1983
4.4517 | -81.1981
4.452 | -81.199
4.4523 | -81.1988
4.4526 | -81.1993
4.453 | -81.201
4.4533 | -81.2018
4.4536 | -81.2013
4.4539 | -81.2008
4.4543 | -81.2
4.4546 | -81.1991
4.4549 | -81.199
4.4552 | -81.1975
4.4555 | -81.1959
4.4559 | -81.1952
4.4562 | -81.195
4.4565 | -81.194
4.4568 | -81.1936
4.4571 | -81.1926
4.4575 | -81.1922
4.4578 | -81.1911
4.4581 | -81.1898
4.4584 | -81.1894
4.4588 | -81.1891
4.4591 | -81.1884
4.4594 | -81.1887
4.4597 | -81.1884
4.46 | -81.1881
4.4604 | -81.188
4.4607 | -81.1876
4.461 | -81.1869
4.4613 | -81.1862
4.4616 | -81.1854
4.462 | -81.1846
4.4623 | -81.1845
4.4626 | -81.1854
4.4629 | -81.1845
4.4633 | -81.1841
4.4636 | -81.1837
4.4639 | -81.1826
4.4642 | -81.1815
4.4645 | -81.1813
4.4649 | -81.1816
4.4652 | -81.1815
4.4655 | -81.1817
4.4658 | -81.1812
4.4661 | -81.1808
4.4665 | -81.1803
4.4668 | -81.1795
4.4671 | -81.1787
4.4674 | -81.1785
4.4678 | -81.1784
4.4681 | -81.1788
4.4684 | -81.179
4.4687 | -81.1788
4.469 | -81.1779
4.4694 | -81.177
4.4697 | -81.1776
4.47 | -81.1782
4.4703 | -81.1784
4.4706 | -81.1781
4.471 | -81.1776
4.4713 | -81.177
4.4716 | -81.1754
4.4719 | -81.1743
4.4722 | -81.1736
4.4726 | -81.1735
4.4729 | -81.1734
4.4732 | -81.1748
4.4735 | -81.1759
4.4739 | -81.1761
4.4742 | -81.1765
4.4745 | -81.1766
4.4748 | -81.1765
4.4751 | -81.1759
4.4755 | -81.1742
4.4758 | -81.1738
4.4761 | -81.1737
4.4764 | -81.1729
4.4767 | -81.1721
4.4771 | -81.1719
4.4774 | -81.1729
4.4777 | -81.1734
4.478 | -81.1746
4.4784 | -81.1734
4.4787 | -81.171
4.479 | -81.1721
4.4793 | -81.1723
4.4796 | -81.1721
4.48 | -81.1716
4.4803 | -81.1711
4.4806 | -81.1709
4.4809 | -81.1706
4.4812 | -81.17
4.4816 | -81.1678
4.4819 | -81.1672
4.4822 | -81.1663
4.4825 | -81.1647
4.4829 | -81.1637
4.4832 | -81.163
4.4835 | -81.1632
4.4838 | -81.1627
4.4841 | -81.1622
4.4845 | -81.1616
4.4848 | -81.1613
4.4851 | -81.1609
4.4854 | -81.1605
4.4857 | -81.1602
4.4861 | -81.1583
4.4864 | -81.1566
4.4867 | -81.1548
4.487 | -81.155
4.4874 | -81.1548
4.4877 | -81.155
4.488 | -81.1554
4.4883 | -81.1549
4.4886 | -81.1545
4.489 | -81.1541
4.4893 | -81.1536
4.4896 | -81.1533
4.4899 | -81.1527
4.4902 | -81.153
4.4906 | -81.153
4.4909 | -81.1527
4.4912 | -81.1527
4.4915 | -81.1528
4.4919 | -81.1526
4.4922 | -81.1522
4.4925 | -81.1521
4.4928 | -81.1518
4.4931 | -81.1516
4.4935 | -81.1512
4.4938 | -81.1511
4.4941 | -81.151
4.4944 | -81.1512
4.4947 | -81.1512
4.4951 | -81.1507
4.4954 | -81.1509
4.4957 | -81.151
4.496 | -81.1523
4.4963 | -81.1525
4.4967 | -81.1548
4.497 | -81.1566
4.4973 | -81.158
4.4976 | -81.158
4.498 | -81.1584
4.4983 | -81.1579
4.4986 | -81.1577
4.4989 | -81.1575
4.4992 | -81.1573
4.4996 | -81.1574
4.4999 | -81.1577
4.5002 | -81.1573
4.5005 | -81.1574
4.5008 | -81.1576
4.5012 | -81.158
4.5015 | -81.1576
4.5018 | -81.1583
4.5021 | -81.1583
4.5025 | -81.1582
4.5028 | -81.1586
4.5031 | -81.1586
4.5034 | -81.1592
4.5037 | -81.1586
4.5041 | -81.1579
4.5044 | -81.1581
4.5047 | -81.1592
4.505 | -81.1601
4.5053 | -81.1609
4.5057 | -81.1625
4.506 | -81.1631
4.5063 | -81.1633
4.5066 | -81.1634
4.507 | -81.1641
4.5073 | -81.1645
4.5076 | -81.1647
4.5079 | -81.1643
4.5082 | -81.1647
4.5086 | -81.1649
4.5089 | -81.1646
4.5092 | -81.1652
4.5095 | -81.1659
4.5098 | -81.1661
4.5102 | -81.1664
4.5105 | -81.1675
4.5108 | -81.167
4.5111 | -81.1673
4.5115 | -81.1672
4.5118 | -81.1662
4.5121 | -81.1658
4.5124 | -81.1661
4.5127 | -81.1671
4.5131 | -81.1665
4.5134 | -81.1652
4.5137 | -81.1656
4.514 | -81.1663
4.5143 | -81.1675
4.5147 | -81.168
4.515 | -81.1694
4.5153 | -81.1693
4.5156 | -81.1694
4.516 | -81.1699
4.5163 | -81.1711
4.5166 | -81.1706
4.5169 | -81.1721
4.5172 | -81.1729
4.5176 | -81.1737
4.5179 | -81.1757
4.5182 | -81.1773
4.5185 | -81.1802
4.5188 | -81.1815
4.5192 | -81.1819
4.5195 | -81.1821
4.5198 | -81.1827
4.5201 | -81.1828
4.5204 | -81.185
4.5208 | -81.1855
4.5211 | -81.1859
4.5214 | -81.1864
4.5217 | -81.1879
4.5221 | -81.1901
4.5224 | -81.1907
4.5227 | -81.1916
4.523 | -81.1921
4.5233 | -81.1933
4.5237 | -81.1946
4.524 | -81.1952
4.5243 | -81.1959
4.5246 | -81.1957
4.5249 | -81.1952
4.5253 | -81.1961
4.5256 | -81.1969
4.5259 | -81.1976
4.5262 | -81.1986
4.5266 | -81.2005
4.5269 | -81.2016
4.5272 | -81.2022
4.5275 | -81.2028
4.5278 | -81.2043
4.5282 | -81.205
4.5285 | -81.2058
4.5288 | -81.2067
4.5291 | -81.2072
4.5294 | -81.2087
4.5298 | -81.21
4.5301 | -81.2107
4.5304 | -81.2117
4.5307 | -81.2124
4.5311 | -81.2121
4.5314 | -81.2128
4.5317 | -81.2136
4.532 | -81.2143
4.5323 | -81.2145
4.5327 | -81.2155
4.533 | -81.2172
4.5333 | -81.2191
4.5336 | -81.2202
4.5339 | -81.2199
4.5343 | -81.22
4.5346 | -81.2212
4.5349 | -81.223
4.5352 | -81.224
4.5356 | -81.225
4.5359 | -81.2258
4.5362 | -81.2266
4.5365 | -81.2271
4.5368 | -81.2276
4.5372 | -81.2283
4.5375 | -81.2293
4.5378 | -81.23
4.5381 | -81.2295
4.5384 | -81.2295
4.5388 | -81.2315
4.5391 | -81.2332
4.5394 | -81.2348
4.5397 | -81.2359
4.54 | -81.2374
4.5404 | -81.2395
4.5407 | -81.2402
4.541 | -81.2412
4.5413 | -81.2423
4.5417 | -81.2437
4.542 | -81.2448
4.5423 | -81.2461
4.5426 | -81.2486
4.5429 | -81.2501
4.5433 | -81.2511
4.5436 | -81.2525
4.5439 | -81.2536
4.5442 | -81.2553
4.5445 | -81.2585
4.5449 | -81.2597
4.5452 | -81.2607
4.5455 | -81.2628
4.5458 | -81.2651
4.5462 | -81.2674
4.5465 | -81.268
4.5468 | -81.2687
4.5471 | -81.2704
4.5474 | -81.2716
4.5478 | -81.2727
4.5481 | -81.2747
4.5484 | -81.2767
4.5487 | -81.2776
4.549 | -81.2793
4.5494 | -81.2814
4.5497 | -81.2847
4.55 | -81.2867
4.5503 | -81.2876
4.5507 | -81.2891
4.551 | -81.2904
4.5513 | -81.2919
4.5516 | -81.2935
4.5519 | -81.2963
4.5523 | -81.2973
4.5526 | -81.2985
4.5529 | -81.3002
4.5532 | -81.3024
4.5535 | -81.3039
4.5539 | -81.3054
4.5542 | -81.3067
4.5545 | -81.3071
4.5548 | -81.3085
4.5552 | -81.3107
4.5555 | -81.3135
4.5558 | -81.3158
4.5561 | -81.3183
4.5564 | -81.3206
4.5568 | -81.3221
4.5571 | -81.3251
4.5574 | -81.3288
4.5577 | -81.3304
4.558 | -81.3317
4.5584 | -81.3343
4.5587 | -81.3377
4.559 | -81.3395
4.5593 | -81.3416
4.5597 | -81.3432
4.56 | -81.3446
4.5603 | -81.3459
4.5606 | -81.3473
4.5609 | -81.3491
4.5613 | -81.3521
4.5616 | -81.3529
4.5619 | -81.3542
4.5622 | -81.356
4.5625 | -81.3577
4.5629 | -81.3591
4.5632 | -81.3609
4.5635 | -81.3625
4.5638 | -81.3644
4.5641 | -81.3652
4.5645 | -81.3673
4.5648 | -81.3688
4.5651 | -81.3693
4.5654 | -81.3705
4.5658 | -81.3723
4.5661 | -81.3741
4.5664 | -81.3751
4.5667 | -81.3767
4.567 | -81.3785
4.5674 | -81.3803
4.5677 | -81.3828
4.568 | -81.3844
4.5683 | -81.3865
4.5686 | -81.3898
4.569 | -81.3926
4.5693 | -81.3952
4.5696 | -81.3975
4.5699 | -81.3989
4.5703 | -81.4009
4.5706 | -81.4032
4.5709 | -81.405
4.5712 | -81.4077
4.5715 | -81.4097
4.5719 | -81.4113
4.5722 | -81.4121
4.5725 | -81.4139
4.5728 | -81.4171
4.5731 | -81.4214
4.5735 | -81.4236
4.5738 | -81.4256
4.5741 | -81.4284
4.5744 | -81.4325
4.5748 | -81.4346
4.5751 | -81.4361
4.5754 | -81.4381
4.5757 | -81.4408
4.576 | -81.4432
4.5764 | -81.4452
4.5767 | -81.4458
4.577 | -81.4463
4.5773 | -81.4482
4.5776 | -81.4494
4.578 | -81.4504
4.5783 | -81.4529
4.5786 | -81.4548
4.5789 | -81.4571
4.5793 | -81.4601
4.5796 | -81.4624
4.5799 | -81.4648
4.5802 | -81.4665
4.5805 | -81.4681
4.5809 | -81.4701
4.5812 | -81.4726
4.5815 | -81.4747
4.5818 | -81.4773
4.5821 | -81.48
4.5825 | -81.4825
4.5828 | -81.4837
4.5831 | -81.4846
4.5834 | -81.4852
4.5838 | -81.4864
4.5841 | -81.4879
4.5844 | -81.4896
4.5847 | -81.4918
4.585 | -81.4946
4.5854 | -81.497
4.5857 | -81.4992
4.586 | -81.5012
4.5863 | -81.5033
4.5866 | -81.5052
4.587 | -81.5065
4.5873 | -81.5098
4.5876 | -81.5124
4.5879 | -81.5135
4.5883 | -81.5149
4.5886 | -81.5172
4.5889 | -81.5194
4.5892 | -81.5215
4.5895 | -81.5232
4.5899 | -81.5254
4.5902 | -81.5273
4.5905 | -81.5291
4.5908 | -81.5313
4.5911 | -81.5326
4.5915 | -81.5338
4.5918 | -81.5355
4.5921 | -81.5374
4.5924 | -81.5394
4.5927 | -81.542
4.5931 | -81.5449
4.5934 | -81.5468
4.5937 | -81.5484
4.594 | -81.5511
4.5944 | -81.5542
4.5947 | -81.5559
4.595 | -81.5573
4.5953 | -81.5593
4.5956 | -81.5602
4.596 | -81.5606
4.5963 | -81.562
4.5966 | -81.5643
4.5969 | -81.5665
4.5972 | -81.5681
4.5976 | -81.5708
4.5979 | -81.5731
4.5982 | -81.5756
4.5985 | -81.5771
4.5989 | -81.5791
4.5992 | -81.5818
4.5995 | -81.5841
4.5998 | -81.5874
4.6001 | -81.5903
4.6005 | -81.5921
4.6008 | -81.5944
4.6011 | -81.5963
4.6014 | -81.5993
4.6017 | -81.6033
4.6021 | -81.6065
4.6024 | -81.6081
4.6027 | -81.6107
4.603 | -81.6131
4.6034 | -81.6153
4.6037 | -81.6183
4.604 | -81.6205
4.6043 | -81.6225
4.6046 | -81.6257
4.605 | -81.6278
4.6053 | -81.6303
4.6056 | -81.6322
4.6059 | -81.634
4.6062 | -81.6362
4.6066 | -81.6384
4.6069 | -81.6409
4.6072 | -81.6434
4.6075 | -81.6458
4.6079 | -81.6496
4.6082 | -81.6537
4.6085 | -81.6558
4.6088 | -81.6577
4.6091 | -81.6595
4.6095 | -81.661
4.6098 | -81.6639
4.6101 | -81.6661
4.6104 | -81.6677
4.6107 | -81.6704
4.6111 | -81.6733
4.6114 | -81.6755
4.6117 | -81.6784
4.612 | -81.6812
4.6123 | -81.6837
4.6127 | -81.6859
4.613 | -81.6875
4.6133 | -81.6894
4.6136 | -81.6913
4.614 | -81.6938
4.6143 | -81.6972
4.6146 | -81.7
4.6149 | -81.7035
4.6152 | -81.7063
4.6156 | -81.7087
4.6159 | -81.7109
4.6162 | -81.7126
4.6165 | -81.7147
4.6168 | -81.7167
4.6172 | -81.7189
4.6175 | -81.7208
4.6178 | -81.7235
4.6181 | -81.7262
4.6185 | -81.7283
4.6188 | -81.7296
4.6191 | -81.7317
4.6194 | -81.7344
4.6197 | -81.7365
4.6201 | -81.7407
4.6204 | -81.7448
4.6207 | -81.7482
4.621 | -81.7506
4.6213 | -81.7524
4.6217 | -81.7559
4.622 | -81.7583
4.6223 | -81.7596
4.6226 | -81.7617
4.623 | -81.7639
4.6233 | -81.7651
4.6236 | -81.767
4.6239 | -81.769
4.6242 | -81.7701
4.6246 | -81.7722
4.6249 | -81.7742
4.6252 | -81.777
4.6255 | -81.7796
4.6258 | -81.7809
4.6262 | -81.7806
4.6265 | -81.7828
4.6268 | -81.7846
4.6271 | -81.7867
4.6275 | -81.7899
4.6278 | -81.7929
4.6281 | -81.7944
4.6284 | -81.796
4.6287 | -81.7978
4.6291 | -81.7994
4.6294 | -81.8014
4.6297 | -81.8034
4.63 | -81.8048
4.6303 | -81.8066
4.6307 | -81.809
4.631 | -81.8111
4.6313 | -81.8131
4.6316 | -81.8137
4.6319 | -81.8155
4.6323 | -81.8171
4.6326 | -81.8195
4.6329 | -81.8224
4.6332 | -81.825
4.6336 | -81.8277
4.6339 | -81.8298
4.6342 | -81.8322
4.6345 | -81.8344
4.6348 | -81.8359
4.6352 | -81.8379
4.6355 | -81.8398
4.6358 | -81.8416
4.6361 | -81.8444
4.6364 | -81.8473
4.6368 | -81.849
4.6371 | -81.8502
4.6374 | -81.8521
4.6377 | -81.8542
4.6381 | -81.8557
4.6384 | -81.8575
4.6387 | -81.86
4.639 | -81.8619
4.6393 | -81.8634
4.6397 | -81.8653
4.64 | -81.8667
4.6403 | -81.8675
4.6406 | -81.8704
4.6409 | -81.8747
4.6413 | -81.8778
4.6416 | -81.8798
4.6419 | -81.8823
4.6422 | -81.8841
4.6426 | -81.8853
4.6429 | -81.8868
4.6432 | -81.8884
4.6435 | -81.8903
4.6438 | -81.8917
4.6442 | -81.8928
4.6445 | -81.8938
4.6448 | -81.8949
4.6451 | -81.8954
4.6454 | -81.8974
4.6458 | -81.8987
4.6461 | -81.9
4.6464 | -81.9014
4.6467 | -81.9023
4.6471 | -81.9039
4.6474 | -81.9054
4.6477 | -81.9073
4.648 | -81.9093
4.6483 | -81.9106
4.6487 | -81.9121
4.649 | -81.9134
4.6493 | -81.9143
4.6496 | -81.9158
4.6499 | -81.9172
4.6503 | -81.9188
4.6506 | -81.9207
4.6509 | -81.9216
4.6512 | -81.9225
4.6516 | -81.9241
4.6519 | -81.9254
4.6522 | -81.9271
4.6525 | -81.9286
4.6528 | -81.9305
4.6532 | -81.9325
4.6535 | -81.9345
4.6538 | -81.9349
4.6541 | -81.9364
4.6544 | -81.9377
4.6548 | -81.9387
4.6551 | -81.9394
4.6554 | -81.9404
4.6557 | -81.9422
4.656 | -81.9434
4.6564 | -81.9424
4.6567 | -81.9433
4.657 | -81.9444
4.6573 | -81.9449
4.6577 | -81.9458
4.658 | -81.9471
4.6583 | -81.9485
4.6586 | -81.9499
4.6589 | -81.9511
4.6593 | -81.9524
4.6596 | -81.9535
4.6599 | -81.9559
4.6602 | -81.9566
4.6605 | -81.957
4.6609 | -81.9578
4.6612 | -81.9591
4.6615 | -81.9593
4.6618 | -81.9591
4.6622 | -81.9599
4.6625 | -81.9606
4.6628 | -81.9618
4.6631 | -81.9622
4.6634 | -81.9644
4.6638 | -81.9655
4.6641 | -81.9662
4.6644 | -81.967
4.6647 | -81.9686
4.665 | -81.9694
4.6654 | -81.9703
4.6657 | -81.9713
4.666 | -81.9707
4.6663 | -81.9706
4.6667 | -81.9713
4.667 | -81.9713
4.6673 | -81.9721
4.6676 | -81.9732
4.6679 | -81.9743
4.6683 | -81.9753
4.6686 | -81.9765
4.6689 | -81.9769
4.6692 | -81.9768
4.6695 | -81.977
4.6699 | -81.9772
4.6702 | -81.9787
4.6705 | -81.9808
4.6708 | -81.9816
4.6712 | -81.983
4.6715 | -81.9842
4.6718 | -81.9837
4.6721 | -81.9845
4.6724 | -81.983
4.6728 | -81.9835
4.6731 | -81.9841
4.6734 | -81.9848
4.6737 | -81.9856
4.674 | -81.9876
4.6744 | -81.9904
4.6747 | -81.9911
4.675 | -81.9913
4.6753 | -81.9916
4.6757 | -81.9925
4.676 | -81.9925
4.6763 | -81.993
4.6766 | -81.9932
4.6769 | -81.9938
4.6773 | -81.995
4.6776 | -81.9964
4.6779 | -81.998
4.6782 | -81.9991
4.6785 | -81.9997
4.6789 | -82.0016
4.6792 | -82.0058
4.6795 | -82.0087
4.6798 | -82.0098
4.6802 | -82.0102
4.6805 | -82.0105
4.6808 | -82.0112
4.6811 | -82.0121
4.6814 | -82.0122
4.6818 | -82.0127
4.6821 | -82.0137
4.6824 | -82.0142
4.6827 | -82.0144
4.683 | -82.0156
4.6834 | -82.0157
4.6837 | -82.0167
4.684 | -82.0187
4.6843 | -82.0201
4.6846 | -82.0208
4.685 | -82.0225
4.6853 | -82.0249
4.6856 | -82.0259
4.6859 | -82.0262
4.6863 | -82.026
4.6866 | -82.0262
4.6869 | -82.025
4.6872 | -82.0249
4.6875 | -82.0249
4.6879 | -82.0238
4.6882 | -82.0228
4.6885 | -82.0232
4.6888 | -82.0237
4.6891 | -82.0239
4.6895 | -82.0232
4.6898 | -82.0193
4.6901 | -82.0166
4.6904 | -82.0183
4.6908 | -82.0197
4.6911 | -82.0196
4.6914 | -82.0201
4.6917 | -82.0212
4.692 | -82.0214
4.6924 | -82.0215
4.6927 | -82.0211
4.693 | -82.0201
4.6933 | -82.0198
4.6936 | -82.02
4.694 | -82.0202
4.6943 | -82.0204
4.6946 | -82.0212
4.6949 | -82.023
4.6953 | -82.0242
4.6956 | -82.0244
4.6959 | -82.0245
4.6962 | -82.0242
4.6965 | -82.0238
4.6969 | -82.0234
4.6972 | -82.0235
4.6975 | -82.0239
4.6978 | -82.0244
4.6981 | -82.0241
4.6985 | -82.0242
4.6988 | -82.0251
4.6991 | -82.0251
4.6994 | -82.0251
4.6998 | -82.0251
4.7001 | -82.0261
4.7004 | -82.0259
4.7007 | -82.0266
4.701 | -82.0269
4.7014 | -82.0279
4.7017 | -82.0286
4.702 | -82.0284
4.7023 | -82.0283
4.7026 | -82.0282
4.703 | -82.0282
4.7033 | -82.0288
4.7036 | -82.029
4.7039 | -82.0283
4.7042 | -82.0279
4.7046 | -82.0277
4.7049 | -82.0284
4.7052 | -82.0278
4.7055 | -82.0273
4.7059 | -82.0274
4.7062 | -82.027
4.7065 | -82.0262
4.7068 | -82.0263
4.7071 | -82.026
4.7075 | -82.0256
4.7078 | -82.0255
4.7081 | -82.0263
4.7084 | -82.0265
4.7087 | -82.0264
4.7091 | -82.0257
4.7094 | -82.0253
4.7097 | -82.0252
4.71 | -82.0247
4.7104 | -82.0241
4.7107 | -82.0237
4.711 | -82.0232
4.7113 | -82.0229
4.7116 | -82.0223
4.712 | -82.0217
4.7123 | -82.0217
4.7126 | -82.023
4.7129 | -82.0238
4.7132 | -82.0232
4.7136 | -82.0228
4.7139 | -82.0223
4.7142 | -82.0222
4.7145 | -82.0224
4.7149 | -82.0226
4.7152 | -82.0236
4.7155 | -82.0263
4.7158 | -82.0269
4.7161 | -82.0265
4.7165 | -82.027
4.7168 | -82.027
4.7171 | -82.0234
4.7174 | -82.021
4.7177 | -82.021
4.7181 | -82.0213
4.7184 | -82.0226
4.7187 | -82.0225
4.719 | -82.0226
4.7194 | -82.022
4.7197 | -82.0212
4.72 | -82.0207
4.7203 | -82.0205
4.7206 | -82.02
4.721 | -82.0194
4.7213 | -82.0192
4.7216 | -82.019
4.7219 | -82.0175
4.7222 | -82.0167
4.7226 | -82.017
4.7229 | -82.0178
4.7232 | -82.0177
4.7235 | -82.0177
4.7238 | -82.0167
4.7242 | -82.0158
4.7245 | -82.0175
4.7248 | -82.0172
4.7251 | -82.016
4.7255 | -82.0152
4.7258 | -82.012
4.7261 | -82.0106
4.7264 | -82.0104
4.7267 | -82.0093
4.7271 | -82.0076
4.7274 | -82.0074
4.7277 | -82.007
4.728 | -82.0062
4.7283 | -82.0056
4.7287 | -82.0051
4.729 | -82.0051
4.7293 | -82.0055
4.7296 | -82.0053
4.73 | -82.0047
4.7303 | -82.0044
4.7306 | -82.0038
4.7309 | -82.0043
4.7312 | -82.0038
4.7316 | -82.0045
4.7319 | -82.0076
4.7322 | -82.0101
4.7325 | -82.0102
4.7328 | -82.0094
4.7332 | -82.0076
4.7335 | -82.0069
4.7338 | -82.0056
4.7341 | -82.0039
4.7345 | -82.0032
4.7348 | -82.0025
4.7351 | -82.0018
4.7354 | -82.0013
4.7357 | -82.0006
4.7361 | -81.9998
4.7364 | -81.9991
4.7367 | -81.9979
4.737 | -81.9966
4.7373 | -81.9962
4.7377 | -81.9949
4.738 | -81.9938
4.7383 | -81.9929
4.7386 | -81.993
4.739 | -81.9914
4.7393 | -81.9906
4.7396 | -81.9899
4.7399 | -81.9876
4.7402 | -81.9855
4.7406 | -81.9863
4.7409 | -81.9868
4.7412 | -81.9859
4.7415 | -81.9855
4.7418 | -81.9853
4.7422 | -81.9845
4.7425 | -81.9841
4.7428 | -81.9835
4.7431 | -81.9839
4.7435 | -81.9837
4.7438 | -81.9833
4.7441 | -81.9818
4.7444 | -81.9799
4.7447 | -81.9759
4.7451 | -81.9728
4.7454 | -81.9707
4.7457 | -81.9689
4.746 | -81.9676
4.7463 | -81.9666
4.7467 | -81.9644
4.747 | -81.9636
4.7473 | -81.9629
4.7476 | -81.9623
4.748 | -81.9618
4.7483 | -81.9613
4.7486 | -81.9606
4.7489 | -81.9595
4.7492 | -81.9586
4.7496 | -81.958
4.7499 | -81.9578
4.7502 | -81.9574
4.7505 | -81.9565
4.7508 | -81.9564
4.7512 | -81.9549
4.7515 | -81.9532
4.7518 | -81.9529
4.7521 | -81.952
4.7524 | -81.9513
4.7528 | -81.9507
4.7531 | -81.9514
4.7534 | -81.9524
4.7537 | -81.9523
4.7541 | -81.9524
4.7544 | -81.9526
4.7547 | -81.9516
4.755 | -81.9507
4.7553 | -81.9492
4.7557 | -81.9486
4.756 | -81.9483
4.7563 | -81.9493
4.7566 | -81.9488
4.7569 | -81.9464
4.7573 | -81.9449
4.7576 | -81.9441
4.7579 | -81.9431
4.7582 | -81.942
4.7586 | -81.9409
4.7589 | -81.9396
4.7592 | -81.9382
4.7595 | -81.9367
4.7598 | -81.9347
4.7602 | -81.9325
4.7605 | -81.9318
4.7608 | -81.9312
4.7611 | -81.929
4.7614 | -81.9276
4.7618 | -81.9265
4.7621 | -81.9256
4.7624 | -81.9243
4.7627 | -81.9228
4.7631 | -81.922
4.7634 | -81.9212
4.7637 | -81.9202
4.764 | -81.9189
4.7643 | -81.9177
4.7647 | -81.9163
4.765 | -81.9144
4.7653 | -81.9128
4.7656 | -81.9119
4.7659 | -81.9123
4.7663 | -81.9119
4.7666 | -81.9107
4.7669 | -81.91
4.7672 | -81.9082
4.7676 | -81.9065
4.7679 | -81.9049
4.7682 | -81.9036
4.7685 | -81.9027
4.7688 | -81.9015
4.7692 | -81.9004
4.7695 | -81.899
4.7698 | -81.898
4.7701 | -81.8984
4.7704 | -81.8977
4.7708 | -81.8979
4.7711 | -81.8978
4.7714 | -81.8968
4.7717 | -81.896
4.7721 | -81.8949
4.7724 | -81.8932
4.7727 | -81.892
4.773 | -81.8912
4.7733 | -81.8898
4.7737 | -81.8873
4.774 | -81.8867
4.7743 | -81.8861
4.7746 | -81.8858
4.7749 | -81.8852
4.7753 | -81.8843
4.7756 | -81.8831
4.7759 | -81.883
4.7762 | -81.8828
4.7765 | -81.8823
4.7769 | -81.8814
4.7772 | -81.8804
4.7775 | -81.8788
4.7778 | -81.8785
4.7782 | -81.8777
4.7785 | -81.8771
4.7788 | -81.8761
4.7791 | -81.8741
4.7794 | -81.8722
4.7798 | -81.8711
4.7801 | -81.8692
4.7804 | -81.8675
4.7807 | -81.8666
4.781 | -81.8663
4.7814 | -81.8658
4.7817 | -81.864
4.782 | -81.8629
4.7823 | -81.8625
4.7827 | -81.8624
4.783 | -81.8624
4.7833 | -81.8615
4.7836 | -81.8615
4.7839 | -81.8604
4.7843 | -81.8599
4.7846 | -81.8597
4.7849 | -81.859
4.7852 | -81.8583
4.7855 | -81.8576
4.7859 | -81.8568
4.7862 | -81.8556
4.7865 | -81.8546
4.7868 | -81.8537
4.7872 | -81.8529
4.7875 | -81.8522
4.7878 | -81.8512
4.7881 | -81.8505
4.7884 | -81.8498
4.7888 | -81.8492
4.7891 | -81.8485
4.7894 | -81.8465
4.7897 | -81.8442
4.79 | -81.8429
4.7904 | -81.8421
4.7907 | -81.8413
4.791 | -81.841
4.7913 | -81.8401
4.7917 | -81.8391
4.792 | -81.8385
4.7923 | -81.8373
4.7926 | -81.8354
4.7929 | -81.8331
4.7933 | -81.8326
4.7936 | -81.8333
4.7939 | -81.8334
4.7942 | -81.8325
4.7945 | -81.8303
4.7949 | -81.8286
4.7952 | -81.8284
4.7955 | -81.8279
4.7958 | -81.8272
4.7961 | -81.8264
4.7965 | -81.8257
4.7968 | -81.825
4.7971 | -81.8234
4.7974 | -81.8236
4.7978 | -81.8232
4.7981 | -81.8224
4.7984 | -81.8218
4.7987 | -81.8212
4.799 | -81.8205
4.7994 | -81.8202
4.7997 | -81.8194
4.8 | -81.8186
4.8003 | -81.8178
4.8006 | -81.8173
4.801 | -81.8187
4.8013 | -81.8181
4.8016 | -81.8178
4.8019 | -81.8169
4.8023 | -81.8156
4.8026 | -81.8145
4.8029 | -81.8143
4.8032 | -81.8146
4.8035 | -81.8136
4.8039 | -81.8125
4.8042 | -81.8134
4.8045 | -81.8127
4.8048 | -81.8127
4.8051 | -81.812
4.8055 | -81.8098
4.8058 | -81.8086
4.8061 | -81.8076
4.8064 | -81.8065
4.8068 | -81.8056
4.8071 | -81.8051
4.8074 | -81.8043
4.8077 | -81.8033
4.808 | -81.8031
4.8084 | -81.8032
4.8087 | -81.8039
4.809 | -81.8049
4.8093 | -81.8058
4.8096 | -81.8058
4.81 | -81.805
4.8103 | -81.8041
4.8106 | -81.8031
4.8109 | -81.8012
4.8113 | -81.8006
4.8116 | -81.8004
4.8119 | -81.7982
4.8122 | -81.7973
4.8125 | -81.7965
4.8129 | -81.7953
4.8132 | -81.7935
4.8135 | -81.79
4.8138 | -81.7875
4.8141 | -81.7868
4.8145 | -81.7858
4.8148 | -81.7845
4.8151 | -81.7835
4.8154 | -81.783
4.8158 | -81.7823
4.8161 | -81.7812
4.8164 | -81.7796
4.8167 | -81.7786
4.817 | -81.777
4.8174 | -81.7747
4.8177 | -81.7736
4.818 | -81.7727
4.8183 | -81.7718
4.8186 | -81.7718
4.819 | -81.7711
4.8193 | -81.7705
4.8196 | -81.7698
4.8199 | -81.7692
4.8202 | -81.7686
4.8206 | -81.7679
4.8209 | -81.7675
4.8212 | -81.7665
4.8215 | -81.7661
4.8219 | -81.7649
4.8222 | -81.7643
4.8225 | -81.7637
4.8228 | -81.7626
4.8231 | -81.7615
4.8235 | -81.7608
4.8238 | -81.76
4.8241 | -81.7594
4.8244 | -81.7583
4.8247 | -81.7569
4.8251 | -81.7559
4.8254 | -81.755
4.8257 | -81.7539
4.826 | -81.7543
4.8264 | -81.7538
4.8267 | -81.7529
4.827 | -81.7519
4.8273 | -81.7507
4.8276 | -81.7506
4.828 | -81.7505
4.8283 | -81.7497
4.8286 | -81.7488
4.8289 | -81.7474
4.8292 | -81.7463
4.8296 | -81.7452
4.8299 | -81.7445
4.8302 | -81.7446
4.8305 | -81.7438
4.8309 | -81.7428
4.8312 | -81.7412
4.8315 | -81.7397
4.8318 | -81.7388
4.8321 | -81.7361
4.8325 | -81.7356
4.8328 | -81.7354
4.8331 | -81.7349
4.8334 | -81.7347
4.8337 | -81.734
4.8341 | -81.7335
4.8344 | -81.733
4.8347 | -81.7325
4.835 | -81.7317
4.8354 | -81.7321
4.8357 | -81.7317
4.836 | -81.7313
4.8363 | -81.7317
4.8366 | -81.7315
4.837 | -81.7301
4.8373 | -81.7265
4.8376 | -81.7232
4.8379 | -81.7217
4.8382 | -81.7206
4.8386 | -81.7192
4.8389 | -81.7174
4.8392 | -81.7158
4.8395 | -81.7154
4.8399 | -81.7152
4.8402 | -81.7145
4.8405 | -81.7144
4.8408 | -81.7141
4.8411 | -81.7133
4.8415 | -81.7126
4.8418 | -81.713
4.8421 | -81.7126
4.8424 | -81.7122
4.8427 | -81.7116
4.8431 | -81.7109
4.8434 | -81.7096
4.8437 | -81.7074
4.844 | -81.7056
4.8443 | -81.7046
4.8447 | -81.7042
4.845 | -81.7036
4.8453 | -81.7035
4.8456 | -81.7016
4.846 | -81.7014
4.8463 | -81.701
4.8466 | -81.7001
4.8469 | -81.6993
4.8472 | -81.6985
4.8476 | -81.6978
4.8479 | -81.6973
4.8482 | -81.6965
4.8485 | -81.6962
4.8488 | -81.6956
4.8492 | -81.6948
4.8495 | -81.6944
4.8498 | -81.6945
4.8501 | -81.6939
4.8505 | -81.6935
4.8508 | -81.6931
4.8511 | -81.6929
4.8514 | -81.6921
4.8517 | -81.6914
4.8521 | -81.6908
4.8524 | -81.6902
4.8527 | -81.6897
4.853 | -81.6892
4.8533 | -81.6886
4.8537 | -81.688
4.854 | -81.6871
4.8543 | -81.6862
4.8546 | -81.6857
4.855 | -81.6845
4.8553 | -81.6831
4.8556 | -81.6826
4.8559 | -81.682
4.8562 | -81.6813
4.8566 | -81.6809
4.8569 | -81.6813
4.8572 | -81.6806
4.8575 | -81.6801
4.8578 | -81.6803
4.8582 | -81.68
4.8585 | -81.6797
4.8588 | -81.679
4.8591 | -81.6785
4.8595 | -81.6778
4.8598 | -81.6763
4.8601 | -81.6745
4.8604 | -81.6751
4.8607 | -81.6763
4.8611 | -81.6759
4.8614 | -81.675
4.8617 | -81.6737
4.862 | -81.6729
4.8623 | -81.672
4.8627 | -81.6701
4.863 | -81.6684
4.8633 | -81.6682
4.8636 | -81.6678
4.8639 | -81.6672
4.8643 | -81.6669
4.8646 | -81.667
4.8649 | -81.6665
4.8652 | -81.6655
4.8656 | -81.6642
4.8659 | -81.6627
4.8662 | -81.6618
4.8665 | -81.661
4.8668 | -81.6604
4.8672 | -81.6602
4.8675 | -81.6595
4.8678 | -81.6581
4.8681 | -81.6567
4.8684 | -81.6553
4.8688 | -81.6541
4.8691 | -81.6533
4.8694 | -81.6523
4.8697 | -81.652
4.8701 | -81.652
4.8704 | -81.6519
4.8707 | -81.6511
4.871 | -81.6491
4.8713 | -81.6478
4.8717 | -81.6468
4.872 | -81.6467
4.8723 | -81.6469
4.8726 | -81.6463
4.8729 | -81.6459
4.8733 | -81.6454
4.8736 | -81.6458
4.8739 | -81.6458
4.8742 | -81.6444
4.8746 | -81.6428
4.8749 | -81.6422
4.8752 | -81.64
4.8755 | -81.6388
4.8758 | -81.6388
4.8762 | -81.6389
4.8765 | -81.6388
4.8768 | -81.6384
4.8771 | -81.6379
4.8774 | -81.6374
4.8778 | -81.6365
4.8781 | -81.6358
4.8784 | -81.635
4.8787 | -81.6341
4.8791 | -81.6335
4.8794 | -81.6317
4.8797 | -81.6285
4.88 | -81.6264
4.8803 | -81.6251
4.8807 | -81.6235
4.881 | -81.6222
4.8813 | -81.6216
4.8816 | -81.6212
4.8819 | -81.6213
4.8823 | -81.6217
4.8826 | -81.621
4.8829 | -81.6206
4.8832 | -81.62
4.8836 | -81.6196
4.8839 | -81.619
4.8842 | -81.6185
4.8845 | -81.6181
4.8848 | -81.617
4.8852 | -81.6167
4.8855 | -81.6167
4.8858 | -81.6166
4.8861 | -81.6165
4.8864 | -81.616
4.8868 | -81.6152
4.8871 | -81.6134
4.8874 | -81.6119
4.8877 | -81.6113
4.888 | -81.6113
4.8884 | -81.6112
4.8887 | -81.6097
4.889 | -81.6088
4.8893 | -81.608
4.8897 | -81.6077
4.89 | -81.608
4.8903 | -81.6074
4.8906 | -81.6065
4.8909 | -81.6064
4.8913 | -81.6064
4.8916 | -81.6049
4.8919 | -81.6033
4.8922 | -81.6012
4.8925 | -81.6014
4.8929 | -81.6011
4.8932 | -81.6015
4.8935 | -81.6027
4.8938 | -81.603
4.8942 | -81.6024
4.8945 | -81.6015
4.8948 | -81.6019
4.8951 | -81.6022
4.8954 | -81.6018
4.8958 | -81.6014
4.8961 | -81.6011
4.8964 | -81.6008
4.8967 | -81.6002
4.897 | -81.5986
4.8974 | -81.5976
4.8977 | -81.5972
4.898 | -81.5968
4.8983 | -81.5967
4.8987 | -81.5965
4.899 | -81.5953
4.8993 | -81.5944
4.8996 | -81.5944
4.8999 | -81.5941
4.9003 | -81.594
4.9006 | -81.5932
4.9009 | -81.5928
4.9012 | -81.5935
4.9015 | -81.5928
4.9019 | -81.5923
4.9022 | -81.592
4.9025 | -81.5921
4.9028 | -81.5909
4.9032 | -81.5887
4.9035 | -81.5873
4.9038 | -81.5865
4.9041 | -81.5865
4.9044 | -81.5857
4.9048 | -81.5856
4.9051 | -81.586
4.9054 | -81.5857
4.9057 | -81.5853
4.906 | -81.5861
4.9064 | -81.5872
4.9067 | -81.5869
4.907 | -81.5875
4.9073 | -81.5878
4.9077 | -81.5855
4.908 | -81.5846
4.9083 | -81.5834
4.9086 | -81.5828
4.9089 | -81.5824
4.9093 | -81.5824
4.9096 | -81.5821
4.9099 | -81.5818
4.9102 | -81.5817
4.9105 | -81.5812
4.9109 | -81.5809
4.9112 | -81.5805
4.9115 | -81.58
4.9118 | -81.5797
4.9121 | -81.58
4.9125 | -81.5806
4.9128 | -81.5811
4.9131 | -81.5822
4.9134 | -81.5826
4.9138 | -81.5835
4.9141 | -81.5834
4.9144 | -81.5829
4.9147 | -81.5813
4.915 | -81.5795
4.9154 | -81.5781
4.9157 | -81.5769
4.916 | -81.5766
4.9163 | -81.5757
4.9166 | -81.5751
4.917 | -81.5756
4.9173 | -81.5754
4.9176 | -81.5752
4.9179 | -81.5747
4.9183 | -81.5725
4.9186 | -81.569
4.9189 | -81.5678
4.9192 | -81.5679
4.9195 | -81.567
4.9199 | -81.5664
4.9202 | -81.5655
4.9205 | -81.5655
4.9208 | -81.5648
4.9211 | -81.5638
4.9215 | -81.5639
4.9218 | -81.5643
4.9221 | -81.5642
4.9224 | -81.5631
4.9228 | -81.5629
4.9231 | -81.5617
4.9234 | -81.5606
4.9237 | -81.5596
4.924 | -81.5589
4.9244 | -81.5589
4.9247 | -81.5588
4.925 | -81.5581
4.9253 | -81.5562
4.9256 | -81.5552
4.926 | -81.5541
4.9263 | -81.5535
4.9266 | -81.5547
4.9269 | -81.5548
4.9273 | -81.554
4.9276 | -81.5531
4.9279 | -81.5519
4.9282 | -81.5499
4.9285 | -81.548
4.9289 | -81.5469
4.9292 | -81.5468
4.9295 | -81.5471
4.9298 | -81.5469
4.9301 | -81.5459
4.9305 | -81.5452
4.9308 | -81.5444
4.9311 | -81.543
4.9314 | -81.5428
4.9318 | -81.5417
4.9321 | -81.5406
4.9324 | -81.54
4.9327 | -81.5398
4.933 | -81.5395
4.9334 | -81.5394
4.9337 | -81.5391
4.934 | -81.5385
4.9343 | -81.5377
4.9346 | -81.5368
4.935 | -81.5353
4.9353 | -81.535
4.9356 | -81.5354
4.9359 | -81.5345
4.9362 | -81.534
4.9366 | -81.5334
4.9369 | -81.5333
4.9372 | -81.5328
4.9375 | -81.5319
4.9379 | -81.5314
4.9382 | -81.531
4.9385 | -81.5297
4.9388 | -81.5288
4.9391 | -81.5286
4.9395 | -81.5282
4.9398 | -81.5272
4.9401 | -81.5261
4.9404 | -81.5256
4.9407 | -81.5242
4.9411 | -81.523
4.9414 | -81.5221
4.9417 | -81.5219
4.942 | -81.5214
4.9424 | -81.5211
4.9427 | -81.5206
4.943 | -81.5201
4.9433 | -81.5184
4.9436 | -81.5165
4.944 | -81.5156
4.9443 | -81.5148
4.9446 | -81.5143
4.9449 | -81.5142
4.9452 | -81.5141
4.9456 | -81.5136
4.9459 | -81.5119
4.9462 | -81.5098
4.9465 | -81.508
4.9469 | -81.507
4.9472 | -81.5066
4.9475 | -81.5058
4.9478 | -81.5049
4.9481 | -81.5048
4.9485 | -81.5048
4.9488 | -81.5043
4.9491 | -81.5041
4.9494 | -81.5037
4.9497 | -81.5032
4.9501 | -81.5025
4.9504 | -81.5019
4.9507 | -81.5014
4.951 | -81.5001
4.9514 | -81.4988
4.9517 | -81.4968
4.952 | -81.4954
4.9523 | -81.4949
4.9526 | -81.4934
4.953 | -81.4917
4.9533 | -81.4911
4.9536 | -81.491
4.9539 | -81.4895
4.9542 | -81.4873
4.9546 | -81.4864
4.9549 | -81.4864
4.9552 | -81.4856
4.9555 | -81.4847
4.9558 | -81.4846
4.9562 | -81.4842
4.9565 | -81.484
4.9568 | -81.4842
4.9571 | -81.4833
4.9575 | -81.4831
4.9578 | -81.4825
4.9581 | -81.4814
4.9584 | -81.4815
4.9587 | -81.4826
4.9591 | -81.482
4.9594 | -81.4807
4.9597 | -81.4801
4.96 | -81.4801
4.9603 | -81.4789
4.9607 | -81.4781
4.961 | -81.4776
4.9613 | -81.4775
4.9616 | -81.4773
4.962 | -81.4768
4.9623 | -81.4766
4.9626 | -81.4762
4.9629 | -81.4743
4.9632 | -81.4722
4.9636 | -81.4711
4.9639 | -81.4701
4.9642 | -81.4694
4.9645 | -81.4686
4.9648 | -81.4685
4.9652 | -81.4676
4.9655 | -81.467
4.9658 | -81.4662
4.9661 | -81.4654
4.9665 | -81.4646
4.9668 | -81.464
4.9671 | -81.4635
4.9674 | -81.4626
4.9677 | -81.4611
4.9681 | -81.4589
4.9684 | -81.4579
4.9687 | -81.4571
4.969 | -81.4569
4.9693 | -81.4561
4.9697 | -81.4555
4.97 | -81.4548
4.9703 | -81.4537
4.9706 | -81.4535
4.971 | -81.4541
4.9713 | -81.4537
4.9716 | -81.453
4.9719 | -81.4524
4.9722 | -81.4523
4.9726 | -81.4514
4.9729 | -81.4504
4.9732 | -81.45
4.9735 | -81.4492
4.9738 | -81.449
4.9742 | -81.4495
4.9745 | -81.4486
4.9748 | -81.4475
4.9751 | -81.4468
4.9755 | -81.4464
4.9758 | -81.446
4.9761 | -81.4449
4.9764 | -81.4442
4.9767 | -81.4436
4.9771 | -81.4425
4.9774 | -81.4418
4.9777 | -81.4416
4.978 | -81.4411
4.9783 | -81.4399
4.9787 | -81.44
4.979 | -81.4398
4.9793 | -81.4391
4.9796 | -81.4393
4.9799 | -81.4386
4.9803 | -81.4377
4.9806 | -81.4371
4.9809 | -81.4372
4.9812 | -81.437
4.9816 | -81.4367
4.9819 | -81.4359
4.9822 | -81.4353
4.9825 | -81.4347
4.9828 | -81.4341
4.9832 | -81.4335
4.9835 | -81.4328
4.9838 | -81.4323
4.9841 | -81.4312
4.9844 | -81.4305
4.9848 | -81.4298
4.9851 | -81.4295
4.9854 | -81.4303
4.9857 | -81.4301
4.9861 | -81.4292
4.9864 | -81.4295
4.9867 | -81.4283
4.987 | -81.427
4.9873 | -81.4263
4.9877 | -81.4224
4.988 | -81.4173
4.9883 | -81.4139
4.9886 | -81.4126
4.9889 | -81.4108
4.9893 | -81.4098
4.9896 | -81.4093
4.9899 | -81.4086
4.9902 | -81.4072
4.9906 | -81.4065
4.9909 | -81.4058
4.9912 | -81.4057
4.9915 | -81.407
4.9918 | -81.4061
4.9922 | -81.4056
4.9925 | -81.4055
4.9928 | -81.405
4.9931 | -81.4035
4.9934 | -81.4024
4.9938 | -81.4016
4.9941 | -81.4001
4.9944 | -81.398
4.9947 | -81.3959
4.9951 | -81.3948
4.9954 | -81.3944
4.9957 | -81.3934
4.996 | -81.392
4.9963 | -81.3915
4.9967 | -81.3914
4.997 | -81.3907
4.9973 | -81.3903
4.9976 | -81.3895
4.9979 | -81.3889
4.9983 | -81.3878
4.9986 | -81.3863
4.9989 | -81.3845
4.9992 | -81.384
4.9996 | -81.3836
4.9999 | -81.3824
5.0002 | -81.3813
5.0005 | -81.3807
5.0008 | -81.3802
5.0012 | -81.3793
5.0015 | -81.3785
5.0018 | -81.3795
5.0021 | -81.3801
5.0024 | -81.3796
5.0028 | -81.3791
5.0031 | -81.378
5.0034 | -81.3773
5.0037 | -81.3762
5.004 | -81.3751
5.0044 | -81.3741
5.0047 | -81.3731
5.005 | -81.3729
5.0053 | -81.372
5.0057 | -81.3715
5.006 | -81.3713
5.0063 | -81.3699
5.0066 | -81.3694
5.0069 | -81.3683
5.0073 | -81.3677
5.0076 | -81.3668
5.0079 | -81.3669
5.0082 | -81.3662
5.0085 | -81.3648
5.0089 | -81.3638
5.0092 | -81.3631
5.0095 | -81.3614
5.0098 | -81.3603
5.0102 | -81.359
5.0105 | -81.3582
5.0108 | -81.3576
5.0111 | -81.3566
5.0114 | -81.3555
5.0118 | -81.3552
5.0121 | -81.3536
5.0124 | -81.3523
5.0127 | -81.3513
5.013 | -81.3506
5.0134 | -81.35
5.0137 | -81.3491
5.014 | -81.3478
5.0143 | -81.3468
5.0147 | -81.3461
5.015 | -81.3454
5.0153 | -81.3446
5.0156 | -81.3441
5.0159 | -81.343
5.0163 | -81.342
5.0166 | -81.341
5.0169 | -81.3394
5.0172 | -81.3379
5.0175 | -81.3369
5.0179 | -81.3357
5.0182 | -81.3355
5.0185 | -81.3348
5.0188 | -81.3337
5.0192 | -81.3327
5.0195 | -81.3322
5.0198 | -81.3317
5.0201 | -81.3307
5.0204 | -81.3301
5.0208 | -81.329
5.0211 | -81.3283
5.0214 | -81.3268
5.0217 | -81.3247
5.022 | -81.3236
5.0224 | -81.3229
5.0227 | -81.3223
5.023 | -81.3215
5.0233 | -81.3211
5.0236 | -81.3204
5.024 | -81.3208
5.0243 | -81.3204
5.0246 | -81.3195
5.0249 | -81.3193
5.0253 | -81.3194
5.0256 | -81.3195
5.0259 | -81.3185
5.0262 | -81.3174
5.0265 | -81.317
5.0269 | -81.3168
5.0272 | -81.3187
5.0275 | -81.3195
5.0278 | -81.3188
5.0281 | -81.3179
5.0285 | -81.3167
5.0288 | -81.3152
5.0291 | -81.3147
5.0294 | -81.3142
5.0298 | -81.3134
5.0301 | -81.3126
5.0304 | -81.3121
5.0307 | -81.3116
5.031 | -81.3105
5.0314 | -81.308
5.0317 | -81.3063
5.032 | -81.3054
5.0323 | -81.3043
5.0326 | -81.3029
5.033 | -81.3011
5.0333 | -81.2996
5.0336 | -81.2988
5.0339 | -81.2989
5.0343 | -81.2983
5.0346 | -81.2981
5.0349 | -81.2973
5.0352 | -81.2964
5.0355 | -81.2954
5.0359 | -81.2947
5.0362 | -81.2941
5.0365 | -81.2931
5.0368 | -81.2918
5.0371 | -81.291
5.0375 | -81.2903
5.0378 | -81.2895
5.0381 | -81.2887
5.0384 | -81.2878
5.0388 | -81.2868
5.0391 | -81.2859
5.0394 | -81.2849
5.0397 | -81.284
5.04 | -81.2834
5.0404 | -81.2821
5.0407 | -81.281
5.041 | -81.28
5.0413 | -81.2786
5.0416 | -81.278
5.042 | -81.2771
5.0423 | -81.2763
5.0426 | -81.2756
5.0429 | -81.2751
5.0433 | -81.2754
5.0436 | -81.2749
5.0439 | -81.2746
5.0442 | -81.2742
5.0445 | -81.2738
5.0449 | -81.2741
5.0452 | -81.2736
5.0455 | -81.2716
5.0458 | -81.2708
5.0461 | -81.2704
5.0465 | -81.269
5.0468 | -81.2683
5.0471 | -81.2675
5.0474 | -81.2672
5.0477 | -81.2669
5.0481 | -81.2663
5.0484 | -81.2655
5.0487 | -81.2652
5.049 | -81.2651
5.0494 | -81.2649
5.0497 | -81.2645
5.05 | -81.2645
5.0503 | -81.264
5.0506 | -81.2631
5.051 | -81.2626
5.0513 | -81.2618
5.0516 | -81.2611
5.0519 | -81.2608
5.0522 | -81.2604
5.0526 | -81.2604
5.0529 | -81.2601
5.0532 | -81.2595
5.0535 | -81.259
5.0539 | -81.2582
5.0542 | -81.2569
5.0545 | -81.2563
5.0548 | -81.2557
5.0551 | -81.2546
5.0555 | -81.2528
5.0558 | -81.2519
5.0561 | -81.2521
5.0564 | -81.2543
5.0567 | -81.2555
5.0571 | -81.2544
5.0574 | -81.254
5.0577 | -81.2528
5.058 | -81.2523
5.0584 | -81.2519
5.0587 | -81.2513
5.059 | -81.2509
5.0593 | -81.2507
5.0596 | -81.2504
5.06 | -81.2502
5.0603 | -81.2507
5.0606 | -81.2492
5.0609 | -81.2469
5.0612 | -81.2442
5.0616 | -81.2422
5.0619 | -81.2418
5.0622 | -81.2416
5.0625 | -81.2404
5.0629 | -81.2398
5.0632 | -81.2388
5.0635 | -81.2376
5.0638 | -81.2374
5.0641 | -81.2373
5.0645 | -81.2366
5.0648 | -81.2372
5.0651 | -81.2376
5.0654 | -81.2375
5.0657 | -81.2392
5.0661 | -81.2394
5.0664 | -81.2394
5.0667 | -81.2406
5.067 | -81.2404
5.0674 | -81.2387
5.0677 | -81.2379
5.068 | -81.2374
5.0683 | -81.2376
5.0686 | -81.2372
5.069 | -81.2368
5.0693 | -81.2359
5.0696 | -81.2357
5.0699 | -81.2354
5.0702 | -81.2353
5.0706 | -81.2351
5.0709 | -81.2345
5.0712 | -81.234
5.0715 | -81.2336
5.0718 | -81.2334
5.0722 | -81.2329
5.0725 | -81.2326
5.0728 | -81.2323
5.0731 | -81.2325
5.0735 | -81.2327
5.0738 | -81.2337
5.0741 | -81.2333
5.0744 | -81.2333
5.0747 | -81.2337
5.0751 | -81.2334
5.0754 | -81.2328
5.0757 | -81.232
5.076 | -81.232
5.0763 | -81.2315
5.0767 | -81.2311
5.077 | -81.2305
5.0773 | -81.2302
5.0776 | -81.2301
5.078 | -81.2295
5.0783 | -81.2293
5.0786 | -81.2298
5.0789 | -81.2302
5.0792 | -81.2303
5.0796 | -81.231
5.0799 | -81.2315
5.0802 | -81.2315
5.0805 | -81.2326
5.0808 | -81.2323
5.0812 | -81.233
5.0815 | -81.2337
5.0818 | -81.2335
5.0821 | -81.2336
5.0825 | -81.2338
5.0828 | -81.2334
5.0831 | -81.2322
5.0834 | -81.232
5.0837 | -81.2321
5.0841 | -81.2319
5.0844 | -81.2321
5.0847 | -81.2314
5.085 | -81.2308
5.0853 | -81.2306
5.0857 | -81.2311
5.086 | -81.2335
5.0863 | -81.2351
5.0866 | -81.236
5.087 | -81.237
5.0873 | -81.2374
5.0876 | -81.2379
5.0879 | -81.2389
5.0882 | -81.2392
5.0886 | -81.2385
5.0889 | -81.2386
5.0892 | -81.2385
5.0895 | -81.2398
5.0898 | -81.2398
5.0902 | -81.2394
5.0905 | -81.2388
5.0908 | -81.2388
5.0911 | -81.2388
5.0915 | -81.2389
5.0918 | -81.2391
5.0921 | -81.2382
5.0924 | -81.2378
5.0927 | -81.2378
5.0931 | -81.2395
5.0934 | -81.2406
5.0937 | -81.2401
5.094 | -81.2396
5.0943 | -81.2395
5.0947 | -81.2396
5.095 | -81.2396
5.0953 | -81.2395
5.0956 | -81.2373
5.0959 | -81.236
5.0963 | -81.2357
5.0966 | -81.2356
5.0969 | -81.2363
5.0972 | -81.2371
5.0976 | -81.2376
5.0979 | -81.2375
5.0982 | -81.2377
5.0985 | -81.2382
5.0988 | -81.2399
5.0992 | -81.2405
5.0995 | -81.2419
5.0998 | -81.2426
5.1001 | -81.2428
5.1004 | -81.2432
5.1008 | -81.244
5.1011 | -81.2442
5.1014 | -81.2436
5.1017 | -81.2437
5.1021 | -81.2439
5.1024 | -81.2447
5.1027 | -81.2447
5.103 | -81.2443
5.1033 | -81.2448
5.1037 | -81.2457
5.104 | -81.246
5.1043 | -81.2457
5.1046 | -81.2452
5.1049 | -81.2453
5.1053 | -81.2458
5.1056 | -81.2461
5.1059 | -81.2465
5.1062 | -81.2471
5.1066 | -81.247
5.1069 | -81.2461
5.1072 | -81.2466
5.1075 | -81.2478
5.1078 | -81.249
5.1082 | -81.2504
5.1085 | -81.2508
5.1088 | -81.2506
5.1091 | -81.2504
5.1094 | -81.2503
5.1098 | -81.25
5.1101 | -81.2506
5.1104 | -81.252
5.1107 | -81.2536
5.1111 | -81.2549
5.1114 | -81.2553
5.1117 | -81.2554
5.112 | -81.2566
5.1123 | -81.2576
5.1127 | -81.2583
5.113 | -81.2597
5.1133 | -81.2603
5.1136 | -81.2598
5.1139 | -81.258
5.1143 | -81.2573
5.1146 | -81.2574
5.1149 | -81.258
5.1152 | -81.2581
5.1155 | -81.2584
5.1159 | -81.2593
5.1162 | -81.2592
5.1165 | -81.26
5.1168 | -81.2619
5.1172 | -81.2623
5.1175 | -81.2636
5.1178 | -81.2635
5.1181 | -81.2635
5.1184 | -81.264
5.1188 | -81.2645
5.1191 | -81.2657
5.1194 | -81.2667
5.1197 | -81.2674
5.12 | -81.2682
5.1204 | -81.2688
5.1207 | -81.2689
5.121 | -81.2693
5.1213 | -81.2706
5.1217 | -81.2714
5.122 | -81.272
5.1223 | -81.273
5.1226 | -81.2744
5.1229 | -81.2753
5.1233 | -81.2764
5.1236 | -81.2774
5.1239 | -81.2783
5.1242 | -81.2784
5.1245 | -81.2792
5.1249 | -81.2798
5.1252 | -81.2806
5.1255 | -81.2819
5.1258 | -81.2825
5.1262 | -81.2833
5.1265 | -81.2842
5.1268 | -81.2863
5.1271 | -81.287
5.1274 | -81.2879
5.1278 | -81.2887
5.1281 | -81.2889
5.1284 | -81.2896
5.1287 | -81.2902
5.129 | -81.2904
5.1294 | -81.2901
5.1297 | -81.2905
5.13 | -81.2918
5.1303 | -81.2927
5.1307 | -81.2939
5.131 | -81.2947
5.1313 | -81.295
5.1316 | -81.2941
5.1319 | -81.2943
5.1323 | -81.2953
5.1326 | -81.2964
5.1329 | -81.2975
5.1332 | -81.2984
5.1335 | -81.2993
5.1339 | -81.3003
5.1342 | -81.3024
5.1345 | -81.3036
5.1348 | -81.3044
5.1352 | -81.3056
5.1355 | -81.3067
5.1358 | -81.3073
5.1361 | -81.309
5.1364 | -81.31
5.1368 | -81.3107
5.1371 | -81.3121
5.1374 | -81.3138
5.1377 | -81.316
5.138 | -81.3176
5.1384 | -81.3185
5.1387 | -81.3192
5.139 | -81.3196
5.1393 | -81.3209
5.1396 | -81.3232
5.14 | -81.3251
5.1403 | -81.3251
5.1406 | -81.3255
5.1409 | -81.3265
5.1413 | -81.327
5.1416 | -81.3271
5.1419 | -81.328
5.1422 | -81.3296
5.1425 | -81.3309
5.1429 | -81.3319
5.1432 | -81.333
5.1435 | -81.3343
5.1438 | -81.3355
5.1441 | -81.3365
5.1445 | -81.3375
5.1448 | -81.3387
5.1451 | -81.3401
5.1454 | -81.3423
5.1458 | -81.3437
5.1461 | -81.3454
5.1464 | -81.347
5.1467 | -81.3485
5.147 | -81.3499
5.1474 | -81.3507
5.1477 | -81.3518
5.148 | -81.3537
5.1483 | -81.3545
5.1486 | -81.3556
5.149 | -81.3564
5.1493 | -81.3576
5.1496 | -81.3589
5.1499 | -81.3603
5.1503 | -81.3613
5.1506 | -81.3622
5.1509 | -81.364
5.1512 | -81.3663
5.1515 | -81.3676
5.1519 | -81.3683
5.1522 | -81.3685
5.1525 | -81.3688
5.1528 | -81.3685
5.1531 | -81.3689
5.1535 | -81.3704
5.1538 | -81.3717
5.1541 | -81.3735
5.1544 | -81.3747
5.1548 | -81.3754
5.1551 | -81.3761
5.1554 | -81.377
5.1557 | -81.3778
5.156 | -81.3789
5.1564 | -81.38
5.1567 | -81.381
5.157 | -81.3832
5.1573 | -81.3845
5.1576 | -81.3855
5.158 | -81.3868
5.1583 | -81.3886
5.1586 | -81.3881
5.1589 | -81.3871
5.1593 | -81.3876
5.1596 | -81.3896
5.1599 | -81.3927
5.1602 | -81.395
5.1605 | -81.3956
5.1609 | -81.3965
5.1612 | -81.3979
5.1615 | -81.3991
5.1618 | -81.3996
5.1621 | -81.4002
5.1625 | -81.401
5.1628 | -81.4032
5.1631 | -81.4048
5.1634 | -81.4066
5.1637 | -81.4083
5.1641 | -81.4102
5.1644 | -81.4115
5.1647 | -81.4122
5.165 | -81.4135
5.1654 | -81.4143
5.1657 | -81.4156
5.166 | -81.4169
5.1663 | -81.4182
5.1666 | -81.4188
5.167 | -81.4197
5.1673 | -81.4208
5.1676 | -81.422
5.1679 | -81.4234
5.1682 | -81.424
5.1686 | -81.4245
5.1689 | -81.426
5.1692 | -81.4274
5.1695 | -81.4282
5.1699 | -81.4297
5.1702 | -81.4307
5.1705 | -81.4312
5.1708 | -81.4323
5.1711 | -81.4336
5.1715 | -81.4352
5.1718 | -81.437
5.1721 | -81.4384
5.1724 | -81.44
5.1727 | -81.4415
5.1731 | -81.4432
5.1734 | -81.4443
5.1737 | -81.4452
5.174 | -81.4465
5.1744 | -81.4479
5.1747 | -81.4497
5.175 | -81.4509
5.1753 | -81.4521
5.1756 | -81.4532
5.176 | -81.4541
5.1763 | -81.4559
5.1766 | -81.4574
5.1769 | -81.4587
5.1772 | -81.4596
5.1776 | -81.4605
5.1779 | -81.4614
5.1782 | -81.4627
5.1785 | -81.4642
5.1789 | -81.466
5.1792 | -81.4674
5.1795 | -81.4685
5.1798 | -81.4715
5.1801 | -81.4732
5.1805 | -81.4742
5.1808 | -81.4756
5.1811 | -81.4767
5.1814 | -81.4766
5.1817 | -81.4771
5.1821 | -81.4779
5.1824 | -81.4774
5.1827 | -81.4787
5.183 | -81.48
5.1833 | -81.4808
5.1837 | -81.4819
5.184 | -81.4832
5.1843 | -81.4844
5.1846 | -81.4856
5.185 | -81.4878
5.1853 | -81.4915
5.1856 | -81.4957
5.1859 | -81.4991
5.1862 | -81.5013
5.1866 | -81.5029
5.1869 | -81.5049
5.1872 | -81.5072
5.1875 | -81.5084
5.1878 | -81.5095
5.1882 | -81.5105
5.1885 | -81.5116
5.1888 | -81.5127
5.1891 | -81.5139
5.1895 | -81.515
5.1898 | -81.5161
5.1901 | -81.517
5.1904 | -81.5181
5.1907 | -81.519
5.1911 | -81.5198
5.1914 | -81.5206
5.1917 | -81.5214
5.192 | -81.5238
5.1923 | -81.526
5.1927 | -81.5279
5.193 | -81.5293
5.1933 | -81.5303
5.1936 | -81.5312
5.194 | -81.5327
5.1943 | -81.5349
5.1946 | -81.5363
5.1949 | -81.5376
5.1952 | -81.5387
5.1956 | -81.5385
5.1959 | -81.5385
5.1962 | -81.5397
5.1965 | -81.5404
5.1968 | -81.5417
5.1972 | -81.5427
5.1975 | -81.5437
5.1978 | -81.545
5.1981 | -81.546
5.1985 | -81.5474
5.1988 | -81.5483
5.1991 | -81.5482
5.1994 | -81.5494
5.1997 | -81.5511
5.2001 | -81.5523
5.2004 | -81.5537
5.2007 | -81.5546
5.201 | -81.5562
5.2013 | -81.5573
5.2017 | -81.5583
5.202 | -81.5595
5.2023 | -81.5609
5.2026 | -81.5623
5.203 | -81.5635
5.2033 | -81.5651
5.2036 | -81.5668
5.2039 | -81.568
5.2042 | -81.5697
5.2046 | -81.5704
5.2049 | -81.5717
5.2052 | -81.573
5.2055 | -81.574
5.2058 | -81.5752
5.2062 | -81.5764
5.2065 | -81.5775
5.2068 | -81.5785
5.2071 | -81.5796
5.2074 | -81.581
5.2078 | -81.5819
5.2081 | -81.5835
5.2084 | -81.585
5.2087 | -81.5862
5.2091 | -81.5877
5.2094 | -81.5901
5.2097 | -81.5926
5.21 | -81.5937
5.2103 | -81.5946
5.2107 | -81.5957
5.211 | -81.5969
5.2113 | -81.5985
5.2116 | -81.5999
5.2119 | -81.6011
5.2123 | -81.602
5.2126 | -81.604
5.2129 | -81.6072
5.2132 | -81.6082
5.2136 | -81.6092
5.2139 | -81.6105
5.2142 | -81.6114
5.2145 | -81.6133
5.2148 | -81.6156
5.2152 | -81.6172
5.2155 | -81.619
5.2158 | -81.621
5.2161 | -81.6218
5.2164 | -81.6224
5.2168 | -81.6238
5.2171 | -81.6253
5.2174 | -81.6264
5.2177 | -81.6279
5.2181 | -81.6296
5.2184 | -81.6313
5.2187 | -81.6331
5.219 | -81.6339
5.2193 | -81.6334
5.2197 | -81.6341
5.22 | -81.6353
5.2203 | -81.6363
5.2206 | -81.6367
5.2209 | -81.6363
5.2213 | -81.6374
5.2216 | -81.6402
5.2219 | -81.6422
5.2222 | -81.6436
5.2226 | -81.6448
5.2229 | -81.6456
5.2232 | -81.6467
5.2235 | -81.6476
5.2238 | -81.6482
5.2242 | -81.6487
5.2245 | -81.6486
5.2248 | -81.6489
5.2251 | -81.651
5.2254 | -81.6522
5.2258 | -81.6548
5.2261 | -81.6567
5.2264 | -81.6576
5.2267 | -81.6586
5.2271 | -81.6596
5.2274 | -81.6612
5.2277 | -81.662
5.228 | -81.6628
5.2283 | -81.6635
5.2287 | -81.664
5.229 | -81.6657
5.2293 | -81.6671
5.2296 | -81.6685
5.2299 | -81.6702
5.2303 | -81.6714
5.2306 | -81.6732
5.2309 | -81.674
5.2312 | -81.6756
5.2315 | -81.6775
5.2319 | -81.6788
5.2322 | -81.6803
5.2325 | -81.6814
5.2328 | -81.6824
5.2332 | -81.6831
5.2335 | -81.6843
5.2338 | -81.6858
5.2341 | -81.6867
5.2344 | -81.6874
5.2348 | -81.688
5.2351 | -81.6888
5.2354 | -81.6897
5.2357 | -81.6907
5.236 | -81.6924
5.2364 | -81.6931
5.2367 | -81.694
5.237 | -81.6948
5.2373 | -81.6959
5.2377 | -81.697
5.238 | -81.6987
5.2383 | -81.6994
5.2386 | -81.7005
5.2389 | -81.7026
5.2393 | -81.7042
5.2396 | -81.7055
5.2399 | -81.7063
5.2402 | -81.7076
5.2405 | -81.7099
5.2409 | -81.7113
5.2412 | -81.7127
5.2415 | -81.7136
5.2418 | -81.7143
5.2422 | -81.7157
5.2425 | -81.717
5.2428 | -81.7177
5.2431 | -81.7186
5.2434 | -81.7205
5.2438 | -81.7228
5.2441 | -81.7251
5.2444 | -81.7262
5.2447 | -81.7272
5.245 | -81.7283
5.2454 | -81.7295
5.2457 | -81.7306
5.246 | -81.731
5.2463 | -81.7318
5.2467 | -81.7329
5.247 | -81.7339
5.2473 | -81.7353
5.2476 | -81.7366
5.2479 | -81.7379
5.2483 | -81.7386
5.2486 | -81.7394
5.2489 | -81.7403
5.2492 | -81.7418
5.2495 | -81.7437
5.2499 | -81.745
5.2502 | -81.7459
5.2505 | -81.7467
5.2508 | -81.7465
5.2511 | -81.747
5.2515 | -81.7488
5.2518 | -81.7503
5.2521 | -81.7516
5.2524 | -81.7528
5.2528 | -81.7542
5.2531 | -81.7559
5.2534 | -81.7568
5.2537 | -81.7577
5.254 | -81.7593
5.2544 | -81.7604
5.2547 | -81.761
5.255 | -81.7607
5.2553 | -81.7621
5.2556 | -81.7653
5.256 | -81.7673
5.2563 | -81.769
5.2566 | -81.7703
5.2569 | -81.7711
5.2573 | -81.7709
5.2576 | -81.7714
5.2579 | -81.7723
5.2582 | -81.7731
5.2585 | -81.774
5.2589 | -81.7745
5.2592 | -81.7748
5.2595 | -81.7758
5.2598 | -81.7767
5.2601 | -81.7777
5.2605 | -81.7791
5.2608 | -81.782
5.2611 | -81.7837
5.2614 | -81.7857
5.2618 | -81.7868
5.2621 | -81.7874
5.2624 | -81.7879
5.2627 | -81.7892
5.263 | -81.7905
5.2634 | -81.7918
5.2637 | -81.7929
5.264 | -81.7938
5.2643 | -81.7948
5.2646 | -81.796
5.265 | -81.797
5.2653 | -81.798
5.2656 | -81.7993
5.2659 | -81.8007
5.2663 | -81.8022
5.2666 | -81.8038
5.2669 | -81.8047
5.2672 | -81.8059
5.2675 | -81.8069
5.2679 | -81.8086
5.2682 | -81.8096
5.2685 | -81.8118
5.2688 | -81.8144
5.2691 | -81.8166
5.2695 | -81.8185
5.2698 | -81.8196
5.2701 | -81.8205
5.2704 | -81.8217
5.2708 | -81.8229
5.2711 | -81.8232
5.2714 | -81.8235
5.2717 | -81.8247
5.272 | -81.8271
5.2724 | -81.8312
5.2727 | -81.8333
5.273 | -81.8344
5.2733 | -81.8361
5.2736 | -81.8381
5.274 | -81.8395
5.2743 | -81.8394
5.2746 | -81.8403
5.2749 | -81.8413
5.2752 | -81.843
5.2756 | -81.8445
5.2759 | -81.8463
5.2762 | -81.8475
5.2765 | -81.8488
5.2769 | -81.8492
5.2772 | -81.8504
5.2775 | -81.8522
5.2778 | -81.8539
5.2781 | -81.8554
5.2785 | -81.8571
5.2788 | -81.8584
5.2791 | -81.8594
5.2794 | -81.861
5.2797 | -81.8632
5.2801 | -81.8647
5.2804 | -81.8662
5.2807 | -81.8681
5.281 | -81.8683
5.2814 | -81.8685
5.2817 | -81.8695
5.282 | -81.8713
5.2823 | -81.8737
5.2826 | -81.876
5.283 | -81.8785
5.2833 | -81.8799
5.2836 | -81.8808
5.2839 | -81.8819
5.2842 | -81.8839
5.2846 | -81.8852
5.2849 | -81.8866
5.2852 | -81.8879
5.2855 | -81.8892
5.2859 | -81.891
5.2862 | -81.8915
5.2865 | -81.8926
5.2868 | -81.8948
5.2871 | -81.8962
5.2875 | -81.8981
5.2878 | -81.8992
5.2881 | -81.9003
5.2884 | -81.9015
5.2887 | -81.903
5.2891 | -81.9044
5.2894 | -81.906
5.2897 | -81.9072
5.29 | -81.909
5.2904 | -81.9101
5.2907 | -81.9116
5.291 | -81.9126
5.2913 | -81.9138
5.2916 | -81.9146
5.292 | -81.9155
5.2923 | -81.9169
5.2926 | -81.9182
5.2929 | -81.9191
5.2932 | -81.9205
5.2936 | -81.9218
5.2939 | -81.9228
5.2942 | -81.924
5.2945 | -81.925
5.2949 | -81.9258
5.2952 | -81.9273
5.2955 | -81.9294
5.2958 | -81.9307
5.2961 | -81.9325
5.2965 | -81.9355
5.2968 | -81.9376
5.2971 | -81.939
5.2974 | -81.9398
5.2977 | -81.941
5.2981 | -81.9425
5.2984 | -81.9438
5.2987 | -81.9453
5.299 | -81.9481
5.2993 | -81.9506
5.2997 | -81.9522
5.3 | -81.9536
5.3003 | -81.9538
5.3006 | -81.955
5.301 | -81.9564
5.3013 | -81.9581
5.3016 | -81.961
5.3019 | -81.9635
5.3022 | -81.966
5.3026 | -81.9684
5.3029 | -81.9707
5.3032 | -81.9721
5.3035 | -81.9735
5.3038 | -81.9751
5.3042 | -81.9762
5.3045 | -81.9773
5.3048 | -81.9778
5.3051 | -81.979
5.3055 | -81.9803
5.3058 | -81.9816
5.3061 | -81.9822
5.3064 | -81.983
5.3067 | -81.9846
5.3071 | -81.9863
5.3074 | -81.9886
5.3077 | -81.9918
5.308 | -81.995
5.3083 | -81.9971
5.3087 | -81.9991
5.309 | -82.0015
5.3093 | -82.0033
5.3096 | -82.0047
5.31 | -82.006
5.3103 | -82.008
5.3106 | -82.01
5.3109 | -82.0116
5.3112 | -82.0129
5.3116 | -82.014
5.3119 | -82.0151
5.3122 | -82.0166
5.3125 | -82.0179
5.3128 | -82.0191
5.3132 | -82.0201
5.3135 | -82.0215
5.3138 | -82.0229
5.3141 | -82.0241
5.3145 | -82.025
5.3148 | -82.0265
5.3151 | -82.0279
5.3154 | -82.0293
5.3157 | -82.0305
5.3161 | -82.0302
5.3164 | -82.0305
5.3167 | -82.0317
5.317 | -82.0319
5.3173 | -82.0327
5.3177 | -82.034
5.318 | -82.0352
5.3183 | -82.0365
5.3186 | -82.0381
5.319 | -82.0396
5.3193 | -82.041
5.3196 | -82.0424
5.3199 | -82.045
5.3202 | -82.0484
5.3206 | -82.0507
5.3209 | -82.0528
5.3212 | -82.0542
5.3215 | -82.0564
5.3218 | -82.058
5.3222 | -82.06
5.3225 | -82.0621
5.3228 | -82.0637
5.3231 | -82.0647
5.3234 | -82.0662
5.3238 | -82.0679
5.3241 | -82.069
5.3244 | -82.0697
5.3247 | -82.0716
5.3251 | -82.0742
5.3254 | -82.0773
5.3257 | -82.078
5.326 | -82.0792
5.3263 | -82.0806
5.3267 | -82.0821
5.327 | -82.0835
5.3273 | -82.0851
5.3276 | -82.0868
5.3279 | -82.0887
5.3283 | -82.0905
5.3286 | -82.0917
5.3289 | -82.0928
5.3292 | -82.094
5.3296 | -82.0955
5.3299 | -82.097
5.3302 | -82.0988
5.3305 | -82.0995
5.3308 | -82.0996
5.3312 | -82.101
5.3315 | -82.1028
5.3318 | -82.1038
5.3321 | -82.1055
5.3324 | -82.1079
5.3328 | -82.11
5.3331 | -82.1112
5.3334 | -82.1124
5.3337 | -82.1133
5.3341 | -82.1147
5.3344 | -82.1167
5.3347 | -82.1185
5.335 | -82.1209
5.3353 | -82.1224
5.3357 | -82.1235
5.336 | -82.1244
5.3363 | -82.1249
5.3366 | -82.1263
5.3369 | -82.1277
5.3373 | -82.1292
5.3376 | -82.1309
5.3379 | -82.1322
5.3382 | -82.1337
5.3386 | -82.1357
5.3389 | -82.1376
5.3392 | -82.1394
5.3395 | -82.141
5.3398 | -82.1428
5.3402 | -82.1449
5.3405 | -82.1462
5.3408 | -82.1477
5.3411 | -82.1494
5.3414 | -82.1512
5.3418 | -82.1529
5.3421 | -82.1548
5.3424 | -82.157
5.3427 | -82.1591
5.343 | -82.1611
5.3434 | -82.1629
5.3437 | -82.1646
5.344 | -82.1662
5.3443 | -82.1678
5.3447 | -82.1692
5.345 | -82.1706
5.3453 | -82.1728
5.3456 | -82.1739
5.3459 | -82.1745
5.3463 | -82.1768
5.3466 | -82.1811
5.3469 | -82.1874
5.3472 | -82.1915
5.3475 | -82.1942
5.3479 | -82.1973
5.3482 | -82.1996
5.3485 | -82.2012
5.3488 | -82.2033
5.3492 | -82.2058
5.3495 | -82.2091
5.3498 | -82.2108
5.3501 | -82.2124
5.3504 | -82.2133
5.3508 | -82.2153
5.3511 | -82.2183
5.3514 | -82.2195
5.3517 | -82.2186
5.352 | -82.2179
5.3524 | -82.2194
5.3527 | -82.2213
5.353 | -82.2229
5.3533 | -82.2248
5.3537 | -82.2278
5.354 | -82.231
5.3543 | -82.233
5.3546 | -82.2348
5.3549 | -82.2371
5.3553 | -82.2395
5.3556 | -82.2411
5.3559 | -82.2425
5.3562 | -82.2438
5.3565 | -82.2452
5.3569 | -82.2476
5.3572 | -82.2497
5.3575 | -82.2509
5.3578 | -82.2526
5.3582 | -82.255
5.3585 | -82.2575
5.3588 | -82.2606
5.3591 | -82.2622
5.3594 | -82.2639
5.3598 | -82.2662
5.3601 | -82.268
5.3604 | -82.2694
5.3607 | -82.2707
5.361 | -82.2716
5.3614 | -82.2729
5.3617 | -82.2743
5.362 | -82.2759
5.3623 | -82.277
5.3627 | -82.2786
5.363 | -82.2805
5.3633 | -82.2819
5.3636 | -82.2836
5.3639 | -82.2853
5.3643 | -82.2869
5.3646 | -82.2888
5.3649 | -82.2905
5.3652 | -82.2924
5.3655 | -82.2941
5.3659 | -82.2957
5.3662 | -82.2981
5.3665 | -82.3001
5.3668 | -82.3018
5.3671 | -82.3037
5.3675 | -82.3055
5.3678 | -82.3079
5.3681 | -82.3097
5.3684 | -82.3115
5.3688 | -82.3135
5.3691 | -82.3156
5.3694 | -82.3178
5.3697 | -82.3196
5.37 | -82.322
5.3704 | -82.3237
5.3707 | -82.3254
5.371 | -82.327
5.3713 | -82.3285
5.3716 | -82.3303
5.372 | -82.3325
5.3723 | -82.3342
5.3726 | -82.3365
5.3729 | -82.3392
5.3733 | -82.3412
5.3736 | -82.3429
5.3739 | -82.3448
5.3742 | -82.3465
5.3745 | -82.3485
5.3749 | -82.3503
5.3752 | -82.3519
5.3755 | -82.3536
5.3758 | -82.3554
5.3761 | -82.3573
5.3765 | -82.359
5.3768 | -82.3609
5.3771 | -82.3628
5.3774 | -82.3647
5.3778 | -82.3664
5.3781 | -82.369
5.3784 | -82.3716
5.3787 | -82.373
5.379 | -82.3747
5.3794 | -82.3759
5.3797 | -82.3775
5.38 | -82.3791
5.3803 | -82.3796
5.3806 | -82.3809
5.381 | -82.3833
5.3813 | -82.3855
5.3816 | -82.3875
5.3819 | -82.3887
5.3823 | -82.3902
5.3826 | -82.3926
5.3829 | -82.3959
5.3832 | -82.3975
5.3835 | -82.3982
5.3839 | -82.3998
5.3842 | -82.4015
5.3845 | -82.4024
5.3848 | -82.4032
5.3851 | -82.4049
5.3855 | -82.4064
5.3858 | -82.4078
5.3861 | -82.4095
5.3864 | -82.412
5.3868 | -82.4133
5.3871 | -82.4148
5.3874 | -82.4167
5.3877 | -82.4186
5.388 | -82.4207
5.3884 | -82.4225
5.3887 | -82.4244
5.389 | -82.4266
5.3893 | -82.4287
5.3896 | -82.4299
5.39 | -82.4319
5.3903 | -82.434
5.3906 | -82.4361
5.3909 | -82.4382
5.3913 | -82.4397
5.3916 | -82.441
5.3919 | -82.4428
5.3922 | -82.4439
5.3925 | -82.4463
5.3929 | -82.4483
5.3932 | -82.4517
5.3935 | -82.4543
5.3938 | -82.4562
5.3941 | -82.4581
5.3945 | -82.4605
5.3948 | -82.4633
5.3951 | -82.465
5.3954 | -82.466
5.3957 | -82.4672
5.3961 | -82.4682
5.3964 | -82.4695
5.3967 | -82.4712
5.397 | -82.4729
5.3974 | -82.4749
5.3977 | -82.4768
5.398 | -82.4793
5.3983 | -82.4823
5.3986 | -82.4843
5.399 | -82.4864
5.3993 | -82.4889
5.3996 | -82.4909
5.3999 | -82.4922
5.4002 | -82.4935
5.4006 | -82.4951
5.4009 | -82.4971
5.4012 | -82.4993
5.4015 | -82.5011
5.4019 | -82.5026
5.4022 | -82.5047
5.4025 | -82.5062
5.4028 | -82.5084
5.4031 | -82.5106
5.4035 | -82.5127
5.4038 | -82.5153
5.4041 | -82.5172
5.4044 | -82.5192
5.4047 | -82.5214
5.4051 | -82.5238
5.4054 | -82.5261
5.4057 | -82.5268
5.406 | -82.5268
5.4064 | -82.5285
5.4067 | -82.5304
5.407 | -82.5321
5.4073 | -82.5338
5.4076 | -82.5354
5.408 | -82.5373
5.4083 | -82.5404
5.4086 | -82.544
5.4089 | -82.5464
5.4092 | -82.5485
5.4096 | -82.551
5.4099 | -82.5529
5.4102 | -82.5548
5.4105 | -82.5566
5.4109 | -82.5584
5.4112 | -82.5597
5.4115 | -82.561
5.4118 | -82.5623
5.4121 | -82.5643
5.4125 | -82.566
5.4128 | -82.5677
5.4131 | -82.5697
5.4134 | -82.5718
5.4137 | -82.5741
5.4141 | -82.5759
5.4144 | -82.5781
5.4147 | -82.5802
5.415 | -82.582
5.4153 | -82.5837
5.4157 | -82.5849
5.416 | -82.5855
5.4163 | -82.5872
5.4166 | -82.5886
5.417 | -82.5897
5.4173 | -82.5917
5.4176 | -82.5944
5.4179 | -82.5955
5.4182 | -82.5972
5.4186 | -82.5989
5.4189 | -82.6017
5.4192 | -82.6049
5.4195 | -82.6062
5.4198 | -82.6068
5.4202 | -82.6083
5.4205 | -82.6101
5.4208 | -82.6115
5.4211 | -82.6122
5.4215 | -82.6138
5.4218 | -82.6145
5.4221 | -82.6157
5.4224 | -82.6168
5.4227 | -82.6186
5.4231 | -82.6207
5.4234 | -82.6227
5.4237 | -82.6235
5.424 | -82.6249
5.4243 | -82.6266
5.4247 | -82.6299
5.425 | -82.6333
5.4253 | -82.6373
5.4256 | -82.6418
5.426 | -82.6445
5.4263 | -82.6465
5.4266 | -82.6478
5.4269 | -82.649
5.4272 | -82.6504
5.4276 | -82.653
5.4279 | -82.6554
5.4282 | -82.6575
5.4285 | -82.659
5.4288 | -82.6608
5.4292 | -82.6628
5.4295 | -82.6647
5.4298 | -82.6663
5.4301 | -82.6679
5.4305 | -82.6696
5.4308 | -82.6709
5.4311 | -82.6715
5.4314 | -82.6732
5.4317 | -82.6764
5.4321 | -82.6785
5.4324 | -82.6801
5.4327 | -82.6818
5.433 | -82.6832
5.4333 | -82.6839
5.4337 | -82.6833
5.434 | -82.6843
5.4343 | -82.6856
5.4346 | -82.6869
5.4349 | -82.6885
5.4353 | -82.6899
5.4356 | -82.6919
5.4359 | -82.6941
5.4362 | -82.6962
5.4366 | -82.698
5.4369 | -82.6996
5.4372 | -82.7015
5.4375 | -82.7035
5.4378 | -82.7044
5.4382 | -82.7056
5.4385 | -82.7072
5.4388 | -82.7088
5.4391 | -82.7107
5.4394 | -82.7123
5.4398 | -82.7141
5.4401 | -82.7163
5.4404 | -82.7183
5.4407 | -82.7203
5.4411 | -82.7208
5.4414 | -82.7219
5.4417 | -82.7234
5.442 | -82.7252
5.4423 | -82.7274
5.4427 | -82.73
5.443 | -82.7316
5.4433 | -82.7335
5.4436 | -82.7348
5.4439 | -82.7363
5.4443 | -82.7379
5.4446 | -82.7395
5.4449 | -82.7409
5.4452 | -82.743
5.4456 | -82.744
5.4459 | -82.7453
5.4462 | -82.747
5.4465 | -82.7486
5.4468 | -82.7502
5.4472 | -82.7509
5.4475 | -82.7517
5.4478 | -82.7532
5.4481 | -82.7546
5.4484 | -82.757
5.4488 | -82.7601
5.4491 | -82.7617
5.4494 | -82.763
5.4497 | -82.7641
5.4501 | -82.7656
5.4504 | -82.7678
5.4507 | -82.7704
5.451 | -82.7729
5.4513 | -82.7746
5.4517 | -82.776
5.452 | -82.7769
5.4523 | -82.7784
5.4526 | -82.7798
5.4529 | -82.7807
5.4533 | -82.7821
5.4536 | -82.7835
5.4539 | -82.7847
5.4542 | -82.7862
5.4546 | -82.7883
5.4549 | -82.7895
5.4552 | -82.7907
5.4555 | -82.7919
5.4558 | -82.7933
5.4562 | -82.795
5.4565 | -82.798
5.4568 | -82.8006
5.4571 | -82.8024
5.4574 | -82.8032
5.4578 | -82.8046
5.4581 | -82.8061
5.4584 | -82.8086
5.4587 | -82.8106
5.4591 | -82.8125
5.4594 | -82.8142
5.4597 | -82.8157
5.46 | -82.8178
5.4603 | -82.8202
5.4607 | -82.822
5.461 | -82.8231
5.4613 | -82.8242
5.4616 | -82.826
5.4619 | -82.8278
5.4623 | -82.8322
5.4626 | -82.8356
5.4629 | -82.8375
5.4632 | -82.839
5.4635 | -82.8402
5.4639 | -82.8417
5.4642 | -82.8428
5.4645 | -82.8447
5.4648 | -82.8463
5.4652 | -82.8479
5.4655 | -82.8491
5.4658 | -82.8505
5.4661 | -82.8521
5.4664 | -82.8536
5.4668 | -82.8551
5.4671 | -82.8566
5.4674 | -82.8578
5.4677 | -82.8593
5.468 | -82.8614
5.4684 | -82.8637
5.4687 | -82.866
5.469 | -82.8672
5.4693 | -82.8688
5.4697 | -82.8706
5.47 | -82.8736
5.4703 | -82.8757
5.4706 | -82.8778
5.4709 | -82.8798
5.4713 | -82.8818
5.4716 | -82.8841
5.4719 | -82.8866
5.4722 | -82.8887
5.4725 | -82.8904
5.4729 | -82.8915
5.4732 | -82.8926
5.4735 | -82.894
5.4738 | -82.8962
5.4742 | -82.8997
5.4745 | -82.9015
5.4748 | -82.9027
5.4751 | -82.9036
5.4754 | -82.9046
5.4758 | -82.9052
5.4761 | -82.9065
5.4764 | -82.9086
5.4767 | -82.9108
5.477 | -82.9129
5.4774 | -82.9149
5.4777 | -82.9158
5.478 | -82.9167
5.4783 | -82.9179
5.4787 | -82.9195
5.479 | -82.9212
5.4793 | -82.9232
5.4796 | -82.9248
5.4799 | -82.9261
5.4803 | -82.9272
5.4806 | -82.9287
5.4809 | -82.9298
5.4812 | -82.9309
5.4815 | -82.9335
5.4819 | -82.9378
5.4822 | -82.9409
5.4825 | -82.9423
5.4828 | -82.9422
5.4832 | -82.9438
5.4835 | -82.9465
5.4838 | -82.9473
5.4841 | -82.9483
5.4844 | -82.9486
5.4848 | -82.9496
5.4851 | -82.9512
5.4854 | -82.9527
5.4857 | -82.9541
5.486 | -82.9562
5.4864 | -82.9582
5.4867 | -82.9593
5.487 | -82.9601
5.4873 | -82.9613
5.4876 | -82.9624
5.488 | -82.9639
5.4883 | -82.9655
5.4886 | -82.9665
5.4889 | -82.9678
5.4893 | -82.9697
5.4896 | -82.9709
5.4899 | -82.9702
5.4902 | -82.9706
5.4905 | -82.9717
5.4909 | -82.9728
5.4912 | -82.9741
5.4915 | -82.9764
5.4918 | -82.9781
5.4921 | -82.9792
5.4925 | -82.9809
5.4928 | -82.9828
5.4931 | -82.9836
5.4934 | -82.9846
5.4938 | -82.9859
5.4941 | -82.9876
5.4944 | -82.9891
5.4947 | -82.9902
5.495 | -82.9908
5.4954 | -82.9915
5.4957 | -82.9925
5.496 | -82.9936
5.4963 | -82.9949
5.4966 | -82.996
5.497 | -82.9978
5.4973 | -82.9994
5.4976 | -82.9993
5.4979 | -82.9995
5.4983 | -83.0004
5.4986 | -83.0029
5.4989 | -83.005
5.4992 | -83.0058
5.4995 | -83.0071
5.4999 | -83.0083
5.5002 | -83.0096
5.5005 | -83.0111
5.5008 | -83.0146
5.5011 | -83.0164
5.5015 | -83.0172
5.5018 | -83.0173
5.5021 | -83.0169
5.5024 | -83.0177
5.5028 | -83.019
5.5031 | -83.0215
5.5034 | -83.0227
5.5037 | -83.0238
5.504 | -83.0248
5.5044 | -83.026
5.5047 | -83.0271
5.505 | -83.0283
5.5053 | -83.0301
5.5056 | -83.0308
5.506 | -83.0321
5.5063 | -83.0332
5.5066 | -83.0342
5.5069 | -83.0357
5.5072 | -83.0373
5.5076 | -83.0386
5.5079 | -83.0389
5.5082 | -83.0396
5.5085 | -83.0409
5.5089 | -83.0429
5.5092 | -83.0438
5.5095 | -83.0445
5.5098 | -83.0455
5.5101 | -83.0482
5.5105 | -83.0501
5.5108 | -83.0512
5.5111 | -83.053
5.5114 | -83.0549
5.5117 | -83.056
5.5121 | -83.0576
5.5124 | -83.0586
5.5127 | -83.0593
5.513 | -83.0596
5.5134 | -83.0601
5.5137 | -83.0612
5.514 | -83.0621
5.5143 | -83.0629
5.5146 | -83.064
5.515 | -83.0649
5.5153 | -83.0661
5.5156 | -83.067
5.5159 | -83.0683
5.5162 | -83.0684
5.5166 | -83.0694
5.5169 | -83.0703
5.5172 | -83.0715
5.5175 | -83.0734
5.5179 | -83.0753
5.5182 | -83.0765
5.5185 | -83.0778
5.5188 | -83.0799
5.5191 | -83.0812
5.5195 | -83.0821
5.5198 | -83.083
5.5201 | -83.0839
5.5204 | -83.0849
5.5207 | -83.0863
5.5211 | -83.0879
5.5214 | -83.0889
5.5217 | -83.0906
5.522 | -83.0926
5.5224 | -83.0946
5.5227 | -83.0977
5.523 | -83.1
5.5233 | -83.1017
5.5236 | -83.1036
5.524 | -83.1043
5.5243 | -83.1048
5.5246 | -83.1058
5.5249 | -83.1072
5.5252 | -83.109
5.5256 | -83.1099
5.5259 | -83.112
5.5262 | -83.1141
5.5265 | -83.116
5.5268 | -83.1167
5.5272 | -83.1177
5.5275 | -83.1185
5.5278 | -83.1193
5.5281 | -83.1197
5.5285 | -83.1197
5.5288 | -83.1211
5.5291 | -83.122
5.5294 | -83.1237
5.5297 | -83.1266
5.5301 | -83.1301
5.5304 | -83.1319
5.5307 | -83.132
5.531 | -83.1328
5.5313 | -83.1342
5.5317 | -83.1357
5.532 | -83.1363
5.5323 | -83.1363
5.5326 | -83.1365
5.533 | -83.1377
5.5333 | -83.1391
5.5336 | -83.1407
5.5339 | -83.1411
5.5342 | -83.1412
5.5346 | -83.1421
5.5349 | -83.143
5.5352 | -83.1433
5.5355 | -83.1438
5.5358 | -83.1451
5.5362 | -83.1462
5.5365 | -83.1471
5.5368 | -83.1476
5.5371 | -83.1482
5.5375 | -83.149
5.5378 | -83.1506
5.5381 | -83.1518
5.5384 | -83.1525
5.5387 | -83.1537
5.5391 | -83.1543
5.5394 | -83.1548
5.5397 | -83.1555
5.54 | -83.157
5.5403 | -83.1584
5.5407 | -83.1594
5.541 | -83.1601
5.5413 | -83.1612
5.5416 | -83.1616
5.542 | -83.1625
5.5423 | -83.1643
5.5426 | -83.1652
5.5429 | -83.1656
5.5432 | -83.1662
5.5436 | -83.167
5.5439 | -83.1679
5.5442 | -83.1673
5.5445 | -83.1681
5.5448 | -83.1691
5.5452 | -83.1701
5.5455 | -83.1705
5.5458 | -83.1707
5.5461 | -83.1709
5.5465 | -83.1716
5.5468 | -83.1725
5.5471 | -83.1733
5.5474 | -83.1744
5.5477 | -83.1755
5.5481 | -83.1761
5.5484 | -83.1774
5.5487 | -83.1781
5.549 | -83.1785
5.5493 | -83.1794
5.5497 | -83.1804
5.55 | -83.1817
5.5503 | -83.183
5.5506 | -83.184
5.551 | -83.1849
5.5513 | -83.1856
5.5516 | -83.1868
5.5519 | -83.188
5.5522 | -83.1886
5.5526 | -83.1897
5.5529 | -83.1908
5.5532 | -83.192
5.5535 | -83.1925
5.5538 | -83.1937
5.5542 | -83.1956
5.5545 | -83.1967
5.5548 | -83.1976
5.5551 | -83.2
5.5554 | -83.2017
5.5558 | -83.2012
5.5561 | -83.2015
5.5564 | -83.2028
5.5567 | -83.2035
5.5571 | -83.2037
5.5574 | -83.2042
5.5577 | -83.2052
5.558 | -83.2062
5.5583 | -83.2071
5.5587 | -83.2083
5.559 | -83.2089
5.5593 | -83.2094
5.5596 | -83.2103
5.5599 | -83.2106
5.5603 | -83.2105
5.5606 | -83.2114
5.5609 | -83.2118
5.5612 | -83.2125
5.5616 | -83.2132
5.5619 | -83.2128
5.5622 | -83.2121
5.5625 | -83.2122
5.5628 | -83.213
5.5632 | -83.214
5.5635 | -83.2153
5.5638 | -83.216
5.5641 | -83.2171
5.5644 | -83.2184
5.5648 | -83.2198
5.5651 | -83.2217
5.5654 | -83.2217
5.5657 | -83.2214
5.5661 | -83.2226
5.5664 | -83.2236
5.5667 | -83.2243
5.567 | -83.2253
5.5673 | -83.226
5.5677 | -83.2254
5.568 | -83.2245
5.5683 | -83.2242
5.5686 | -83.2243
5.5689 | -83.2253
5.5693 | -83.2279
5.5696 | -83.2303
5.5699 | -83.2311
5.5702 | -83.2323
5.5706 | -83.2332
5.5709 | -83.234
5.5712 | -83.2355
5.5715 | -83.236
5.5718 | -83.2362
5.5722 | -83.2364
5.5725 | -83.2371
5.5728 | -83.2385
5.5731 | -83.2407
5.5734 | -83.2418
5.5738 | -83.2425
5.5741 | -83.2429
5.5744 | -83.2437
5.5747 | -83.2447
5.575 | -83.2448
5.5754 | -83.2454
5.5757 | -83.2461
5.576 | -83.2469
5.5763 | -83.2475
5.5767 | -83.2476
5.577 | -83.2479
5.5773 | -83.2489
5.5776 | -83.2498
5.5779 | -83.2503
5.5783 | -83.2507
5.5786 | -83.2505
5.5789 | -83.252
5.5792 | -83.2524
5.5795 | -83.2532
5.5799 | -83.2538
5.5802 | -83.2542
5.5805 | -83.255
5.5808 | -83.2561
5.5812 | -83.2571
5.5815 | -83.2579
5.5818 | -83.2587
5.5821 | -83.2591
5.5824 | -83.2591
5.5828 | -83.2597
5.5831 | -83.2602
5.5834 | -83.2612
5.5837 | -83.2619
5.584 | -83.2625
5.5844 | -83.2631
5.5847 | -83.264
5.585 | -83.2649
5.5853 | -83.2661
5.5857 | -83.2667
5.586 | -83.2662
5.5863 | -83.2665
5.5866 | -83.268
5.5869 | -83.269
5.5873 | -83.2705
5.5876 | -83.2715
5.5879 | -83.2706
5.5882 | -83.2699
5.5885 | -83.2706
5.5889 | -83.2714
5.5892 | -83.2722
5.5895 | -83.2726
5.5898 | -83.2725
5.5902 | -83.2719
5.5905 | -83.2722
5.5908 | -83.2727
5.5911 | -83.2731
5.5914 | -83.2745
5.5918 | -83.2759
5.5921 | -83.2765
5.5924 | -83.2769
5.5927 | -83.2777
5.593 | -83.2778
5.5934 | -83.2783
5.5937 | -83.2784
5.594 | -83.279
5.5943 | -83.2807
5.5947 | -83.2827
5.595 | -83.2841
5.5953 | -83.2848
5.5956 | -83.2852
5.5959 | -83.2854
5.5963 | -83.2868
5.5966 | -83.289
5.5969 | -83.2894
5.5972 | -83.2899
5.5975 | -83.2903
5.5979 | -83.2903
5.5982 | -83.2898
5.5985 | -83.2895
5.5988 | -83.2892
5.5991 | -83.2897
5.5995 | -83.2904
5.5998 | -83.2911
5.6001 | -83.2919
5.6004 | -83.2931
5.6008 | -83.2941
5.6011 | -83.2939
5.6014 | -83.2937
5.6017 | -83.2938
5.602 | -83.2936
5.6024 | -83.2926
5.6027 | -83.2929
5.603 | -83.2931
5.6033 | -83.2935
5.6036 | -83.2941
5.604 | -83.2944
5.6043 | -83.2955
5.6046 | -83.2965
5.6049 | -83.2972
5.6053 | -83.298
5.6056 | -83.2984
5.6059 | -83.2987
5.6062 | -83.299
5.6065 | -83.299
5.6069 | -83.2997
5.6072 | -83.3039
5.6075 | -83.3062
5.6078 | -83.3051
5.6081 | -83.305
5.6085 | -83.3052
5.6088 | -83.3044
5.6091 | -83.3043
5.6094 | -83.3044
5.6098 | -83.3048
5.6101 | -83.305
5.6104 | -83.3051
5.6107 | -83.305
5.611 | -83.3053
5.6114 | -83.3061
5.6117 | -83.3065
5.612 | -83.3067
5.6123 | -83.3069
5.6126 | -83.3072
5.613 | -83.3083
5.6133 | -83.3089
5.6136 | -83.3091
5.6139 | -83.3095
5.6143 | -83.3087
5.6146 | -83.3063
5.6149 | -83.3045
5.6152 | -83.3047
5.6155 | -83.305
5.6159 | -83.3043
5.6162 | -83.304
5.6165 | -83.3041
5.6168 | -83.3046
5.6171 | -83.3047
5.6175 | -83.3048
5.6178 | -83.3051
5.6181 | -83.3057
5.6184 | -83.3058
5.6188 | -83.3064
5.6191 | -83.3074
5.6194 | -83.3076
5.6197 | -83.3068
5.62 | -83.307
5.6204 | -83.3072
5.6207 | -83.3074
5.621 | -83.3076
5.6213 | -83.3068
5.6216 | -83.3064
5.622 | -83.3065
5.6223 | -83.3066
5.6226 | -83.3061
5.6229 | -83.3058
5.6232 | -83.3068
5.6236 | -83.3072
5.6239 | -83.3076
5.6242 | -83.3079
5.6245 | -83.3081
5.6249 | -83.3075
5.6252 | -83.3066
5.6255 | -83.3067
5.6258 | -83.3073
5.6261 | -83.3075
5.6265 | -83.3066
5.6268 | -83.3062
5.6271 | -83.3068
5.6274 | -83.3068
5.6277 | -83.3063
5.6281 | -83.3063
5.6284 | -83.3067
5.6287 | -83.3086
5.629 | -83.3094
5.6294 | -83.3091
5.6297 | -83.3091
5.63 | -83.3089
5.6303 | -83.3087
5.6306 | -83.3086
5.631 | -83.3084
5.6313 | -83.3077
5.6316 | -83.3072
5.6319 | -83.307
5.6322 | -83.3068
5.6326 | -83.3073
5.6329 | -83.309
5.6332 | -83.3094
5.6335 | -83.3095
5.6339 | -83.3085
5.6342 | -83.3072
5.6345 | -83.3072
5.6348 | -83.3073
5.6351 | -83.3081
5.6355 | -83.3095
5.6358 | -83.311
5.6361 | -83.3112
5.6364 | -83.3114
5.6367 | -83.3118
5.6371 | -83.3119
5.6374 | -83.3125
5.6377 | -83.3135
5.638 | -83.3134
5.6384 | -83.3135
5.6387 | -83.3129
5.639 | -83.3122
5.6393 | -83.3122
5.6396 | -83.3119
5.64 | -83.311
5.6403 | -83.3104
5.6406 | -83.3109
5.6409 | -83.3114
5.6412 | -83.3111
5.6416 | -83.3105
5.6419 | -83.3103
5.6422 | -83.3105
5.6425 | -83.3103
5.6429 | -83.3104
5.6432 | -83.31
5.6435 | -83.3099
5.6438 | -83.3098
5.6441 | -83.3094
5.6445 | -83.3096
5.6448 | -83.31
5.6451 | -83.3097
5.6454 | -83.3095
5.6457 | -83.3089
5.6461 | -83.308
5.6464 | -83.3082
5.6467 | -83.3088
5.647 | -83.3085
5.6473 | -83.3088
5.6477 | -83.3084
5.648 | -83.3086
5.6483 | -83.3081
5.6486 | -83.3074
5.649 | -83.3069
5.6493 | -83.3071
5.6496 | -83.3074
5.6499 | -83.3071
5.6502 | -83.3078
5.6506 | -83.3082
5.6509 | -83.3077
5.6512 | -83.3091
5.6515 | -83.3087
5.6518 | -83.3082
5.6522 | -83.3077
5.6525 | -83.3066
5.6528 | -83.3065
5.6531 | -83.3067
5.6535 | -83.3079
5.6538 | -83.3086
5.6541 | -83.3087
5.6544 | -83.3084
5.6547 | -83.308
5.6551 | -83.308
5.6554 | -83.3076
5.6557 | -83.3066
5.656 | -83.3062
5.6563 | -83.3064
5.6567 | -83.3061
5.657 | -83.3057
5.6573 | -83.305
5.6576 | -83.3048
5.658 | -83.3042
5.6583 | -83.3042
5.6586 | -83.3022
5.6589 | -83.2999
5.6592 | -83.2991
5.6596 | -83.3
5.6599 | -83.3008
5.6602 | -83.3009
5.6605 | -83.3006
5.6608 | -83.301
5.6612 | -83.3017
5.6615 | -83.3019
5.6618 | -83.3013
5.6621 | -83.301
5.6625 | -83.3004
5.6628 | -83.3002
5.6631 | -83.3001
5.6634 | -83.2996
5.6637 | -83.2988
5.6641 | -83.2981
5.6644 | -83.2974
5.6647 | -83.2969
5.665 | -83.2965
5.6653 | -83.2961
5.6657 | -83.2955
5.666 | -83.2952
5.6663 | -83.2955
5.6666 | -83.2958
5.6669 | -83.2961
5.6673 | -83.2954
5.6676 | -83.2942
5.6679 | -83.2938
5.6682 | -83.2931
5.6686 | -83.2923
5.6689 | -83.2919
5.6692 | -83.2913
5.6695 | -83.2904
5.6698 | -83.2895
5.6702 | -83.2893
5.6705 | -83.2892
5.6708 | -83.2893
5.6711 | -83.2889
5.6714 | -83.2887
5.6718 | -83.288
5.6721 | -83.288
5.6724 | -83.2879
5.6727 | -83.2876
5.6731 | -83.2875
5.6734 | -83.2874
5.6737 | -83.287
5.674 | -83.2862
5.6743 | -83.2851
5.6747 | -83.2858
5.675 | -83.2859
5.6753 | -83.286
5.6756 | -83.2859
5.6759 | -83.2855
5.6763 | -83.2852
5.6766 | -83.2848
5.6769 | -83.2844
5.6772 | -83.285
5.6776 | -83.2847
5.6779 | -83.2842
5.6782 | -83.2837
5.6785 | -83.2836
5.6788 | -83.2825
5.6792 | -83.2798
5.6795 | -83.2752
5.6798 | -83.2728
5.6801 | -83.2722
5.6804 | -83.2711
5.6808 | -83.2688
5.6811 | -83.2673
5.6814 | -83.2681
5.6817 | -83.2672
5.6821 | -83.2667
5.6824 | -83.2661
5.6827 | -83.2651
5.683 | -83.2647
5.6833 | -83.2642
5.6837 | -83.2637
5.684 | -83.2625
5.6843 | -83.2618
5.6846 | -83.2611
5.6849 | -83.2607
5.6853 | -83.2598
5.6856 | -83.2593
5.6859 | -83.2597
5.6862 | -83.2596
5.6866 | -83.2599
5.6869 | -83.2596
5.6872 | -83.2591
5.6875 | -83.259
5.6878 | -83.2601
5.6882 | -83.2604
5.6885 | -83.2596
5.6888 | -83.2585
5.6891 | -83.2578
5.6894 | -83.2565
5.6898 | -83.2553
5.6901 | -83.255
5.6904 | -83.2545
5.6907 | -83.254
5.691 | -83.254
5.6914 | -83.2533
5.6917 | -83.2529
5.692 | -83.253
5.6923 | -83.2525
5.6927 | -83.2529
5.693 | -83.2525
5.6933 | -83.2511
5.6936 | -83.251
5.6939 | -83.2509
5.6943 | -83.2502
5.6946 | -83.2506
5.6949 | -83.2503
5.6952 | -83.2495
5.6955 | -83.2492
5.6959 | -83.2489
5.6962 | -83.2486
5.6965 | -83.2485
5.6968 | -83.2486
5.6972 | -83.2484
5.6975 | -83.2477
5.6978 | -83.2472
5.6981 | -83.247
5.6984 | -83.2493
5.6988 | -83.2496
5.6991 | -83.2481
5.6994 | -83.248
5.6997 | -83.2472
5.7 | -83.2461
5.7004 | -83.2457
5.7007 | -83.245
5.701 | -83.2442
5.7013 | -83.2436
5.7017 | -83.2432
5.702 | -83.2427
5.7023 | -83.2421
5.7026 | -83.2415
5.7029 | -83.2402
5.7033 | -83.2396
5.7036 | -83.2384
5.7039 | -83.2376
5.7042 | -83.2368
5.7045 | -83.2349
5.7049 | -83.2323
5.7052 | -83.2314
5.7055 | -83.2296
5.7058 | -83.2274
5.7062 | -83.227
5.7065 | -83.2262
5.7068 | -83.2253
5.7071 | -83.2253
5.7074 | -83.2253
5.7078 | -83.2241
5.7081 | -83.2216
5.7084 | -83.2205
5.7087 | -83.2208
5.709 | -83.2207
5.7094 | -83.2195
5.7097 | -83.2182
5.71 | -83.2179
5.7103 | -83.2171
5.7107 | -83.2164
5.711 | -83.2141
5.7113 | -83.2136
5.7116 | -83.2125
5.7119 | -83.2123
5.7123 | -83.2116
5.7126 | -83.2112
5.7129 | -83.2105
5.7132 | -83.2094
5.7135 | -83.2074
5.7139 | -83.2075
5.7142 | -83.2085
5.7145 | -83.2088
5.7148 | -83.208
5.7151 | -83.2071
5.7155 | -83.207
5.7158 | -83.2064
5.7161 | -83.2057
5.7164 | -83.2052
5.7168 | -83.2045
5.7171 | -83.2041
5.7174 | -83.2042
5.7177 | -83.2044
5.718 | -83.2034
5.7184 | -83.2023
5.7187 | -83.2015
5.719 | -83.201
5.7193 | -83.2007
5.7196 | -83.2016
5.72 | -83.2011
5.7203 | -83.2025
5.7206 | -83.2024
5.7209 | -83.2019
5.7213 | -83.2012
5.7216 | -83.1989
5.7219 | -83.1947
5.7222 | -83.1924
5.7225 | -83.1917
5.7229 | -83.191
5.7232 | -83.1897
5.7235 | -83.1871
5.7238 | -83.1854
5.7241 | -83.1832
5.7245 | -83.1818
5.7248 | -83.1815
5.7251 | -83.1803
5.7254 | -83.1793
5.7258 | -83.1779
5.7261 | -83.1766
5.7264 | -83.1756
5.7267 | -83.1742
5.727 | -83.1726
5.7274 | -83.1713
5.7277 | -83.1698
5.728 | -83.1682
5.7283 | -83.1671
5.7286 | -83.1654
5.729 | -83.1655
5.7293 | -83.165
5.7296 | -83.1641
5.7299 | -83.1626
5.7303 | -83.1635
5.7306 | -83.1632
5.7309 | -83.1621
5.7312 | -83.1624
5.7315 | -83.1618
5.7319 | -83.1612
5.7322 | -83.1605
5.7325 | -83.1594
5.7328 | -83.1588
5.7331 | -83.158
5.7335 | -83.1574
5.7338 | -83.1565
5.7341 | -83.1552
5.7344 | -83.1536
5.7348 | -83.1527
5.7351 | -83.1519
5.7354 | -83.1518
5.7357 | -83.151
5.736 | -83.1501
5.7364 | -83.15
5.7367 | -83.1518
5.737 | -83.1536
5.7373 | -83.1533
5.7376 | -83.1522
5.738 | -83.1509
5.7383 | -83.1504
5.7386 | -83.1498
5.7389 | -83.1486
5.7392 | -83.1478
5.7396 | -83.147
5.7399 | -83.146
5.7402 | -83.1451
5.7405 | -83.1442
5.7409 | -83.1433
5.7412 | -83.1427
5.7415 | -83.1418
5.7418 | -83.1415
5.7421 | -83.1409
5.7425 | -83.1401
5.7428 | -83.1397
5.7431 | -83.1393
5.7434 | -83.1384
5.7437 | -83.1368
5.7441 | -83.1358
5.7444 | -83.1344
5.7447 | -83.133
5.745 | -83.1322
5.7454 | -83.1314
5.7457 | -83.1312
5.746 | -83.1307
5.7463 | -83.13
5.7466 | -83.1298
5.747 | -83.1292
5.7473 | -83.1284
5.7476 | -83.1261
5.7479 | -83.1246
5.7482 | -83.1242
5.7486 | -83.1237
5.7489 | -83.1228
5.7492 | -83.1198
5.7495 | -83.1186
5.7499 | -83.1179
5.7502 | -83.1158
5.7505 | -83.115
5.7508 | -83.1143
5.7511 | -83.1138
5.7515 | -83.1131
5.7518 | -83.1119
5.7521 | -83.1102
5.7524 | -83.109
5.7527 | -83.1077
5.7531 | -83.1067
5.7534 | -83.1057
5.7537 | -83.1049
5.754 | -83.104
5.7544 | -83.1032
5.7547 | -83.1021
5.755 | -83.1007
5.7553 | -83.1006
5.7556 | -83.1003
5.756 | -83.0994
5.7563 | -83.0986
5.7566 | -83.0983
5.7569 | -83.0973
5.7572 | -83.0971
5.7576 | -83.0964
5.7579 | -83.0955
5.7582 | -83.0948
5.7585 | -83.0939
5.7588 | -83.0931
5.7592 | -83.0925
5.7595 | -83.0918
5.7598 | -83.0911
5.7601 | -83.0903
5.7605 | -83.0893
5.7608 | -83.0882
5.7611 | -83.0873
5.7614 | -83.0863
5.7617 | -83.0851
5.7621 | -83.0837
5.7624 | -83.083
5.7627 | -83.0823
5.763 | -83.0816
5.7633 | -83.0807
5.7637 | -83.0794
5.764 | -83.0772
5.7643 | -83.0746
5.7646 | -83.0725
5.765 | -83.0707
5.7653 | -83.0699
5.7656 | -83.0688
5.7659 | -83.0685
5.7662 | -83.0678
5.7666 | -83.0676
5.7669 | -83.0669
5.7672 | -83.0664
5.7675 | -83.0673
5.7678 | -83.0673
5.7682 | -83.0661
5.7685 | -83.0653
5.7688 | -83.0646
5.7691 | -83.0633
5.7695 | -83.0621
5.7698 | -83.0613
5.7701 | -83.0604
5.7704 | -83.0598
5.7707 | -83.0602
5.7711 | -83.0595
5.7714 | -83.0588
5.7717 | -83.0587
5.772 | -83.0588
5.7723 | -83.0581
5.7727 | -83.0569
5.773 | -83.0546
5.7733 | -83.053
5.7736 | -83.0522
5.774 | -83.0512
5.7743 | -83.0505
5.7746 | -83.0499
5.7749 | -83.0486
5.7752 | -83.0472
5.7756 | -83.0464
5.7759 | -83.046
5.7762 | -83.045
5.7765 | -83.0443
5.7768 | -83.0435
5.7772 | -83.0434
5.7775 | -83.0431
5.7778 | -83.0419
5.7781 | -83.0404
5.7785 | -83.0392
5.7788 | -83.0366
5.7791 | -83.0348
5.7794 | -83.0346
5.7797 | -83.0333
5.7801 | -83.0325
5.7804 | -83.0311
5.7807 | -83.0308
5.781 | -83.0293
5.7813 | -83.0269
5.7817 | -83.024
5.782 | -83.0226
5.7823 | -83.0218
5.7826 | -83.0206
5.7829 | -83.0207
5.7833 | -83.0201
5.7836 | -83.0198
5.7839 | -83.0189
5.7842 | -83.0185
5.7846 | -83.0175
5.7849 | -83.0169
5.7852 | -83.0161
5.7855 | -83.0154
5.7858 | -83.0142
5.7862 | -83.0133
5.7865 | -83.0121
5.7868 | -83.0111
5.7871 | -83.0103
5.7874 | -83.0091
5.7878 | -83.0082
5.7881 | -83.0073
5.7884 | -83.0064
5.7887 | -83.0054
5.7891 | -83.0044
5.7894 | -83.0024
5.7897 | -83.0013
5.79 | -83
5.7903 | -82.9992
5.7907 | -82.9978
5.791 | -82.9971
5.7913 | -82.9964
5.7916 | -82.9954
5.7919 | -82.9942
5.7923 | -82.9922
5.7926 | -82.9912
5.7929 | -82.9901
5.7932 | -82.9887
5.7936 | -82.9879
5.7939 | -82.987
5.7942 | -82.9849
5.7945 | -82.9839
5.7948 | -82.9843
5.7952 | -82.9852
5.7955 | -82.9844
5.7958 | -82.9833
5.7961 | -82.982
5.7964 | -82.9807
5.7968 | -82.9797
5.7971 | -82.9786
5.7974 | -82.9776
5.7977 | -82.9767
5.7981 | -82.9748
5.7984 | -82.9732
5.7987 | -82.9722
5.799 | -82.9714
5.7993 | -82.9705
5.7997 | -82.9697
5.8 | -82.969
5.8003 | -82.9684
5.8006 | -82.9679
5.8009 | -82.9675
5.8013 | -82.9669
5.8016 | -82.9659
5.8019 | -82.9645
5.8022 | -82.9632
5.8026 | -82.9613
5.8029 | -82.9607
5.8032 | -82.9597
5.8035 | -82.9578
5.8038 | -82.9573
5.8042 | -82.9564
5.8045 | -82.955
5.8048 | -82.9538
5.8051 | -82.952
5.8054 | -82.9513
5.8058 | -82.952
5.8061 | -82.9513
5.8064 | -82.9506
5.8067 | -82.9505
5.807 | -82.9499
5.8074 | -82.9492
5.8077 | -82.9497
5.808 | -82.9486
5.8083 | -82.9476
5.8087 | -82.946
5.809 | -82.9452
5.8093 | -82.9452
5.8096 | -82.9456
5.8099 | -82.9433
5.8103 | -82.9417
5.8106 | -82.941
5.8109 | -82.941
5.8112 | -82.9406
5.8115 | -82.9397
5.8119 | -82.939
5.8122 | -82.9385
5.8125 | -82.9377
5.8128 | -82.937
5.8132 | -82.935
5.8135 | -82.9338
5.8138 | -82.933
5.8141 | -82.932
5.8144 | -82.9303
5.8148 | -82.9292
5.8151 | -82.9276
5.8154 | -82.9265
5.8157 | -82.9259
5.816 | -82.9248
5.8164 | -82.9238
5.8167 | -82.9218
5.817 | -82.9214
5.8173 | -82.9217
5.8177 | -82.9211
5.818 | -82.9202
5.8183 | -82.9192
5.8186 | -82.9183
5.8189 | -82.9169
5.8193 | -82.9148
5.8196 | -82.9131
5.8199 | -82.9113
5.8202 | -82.9096
5.8205 | -82.9084
5.8209 | -82.9072
5.8212 | -82.9067
5.8215 | -82.9069
5.8218 | -82.9067
5.8222 | -82.9061
5.8225 | -82.9053
5.8228 | -82.9046
5.8231 | -82.9039
5.8234 | -82.9032
5.8238 | -82.9021
5.8241 | -82.9009
5.8244 | -82.9008
5.8247 | -82.9019
5.825 | -82.902
5.8254 | -82.9013
5.8257 | -82.9
5.826 | -82.899
5.8263 | -82.8981
5.8266 | -82.8971
5.827 | -82.8966
5.8273 | -82.8958
5.8276 | -82.8949
5.8279 | -82.894
5.8283 | -82.8932
5.8286 | -82.8925
5.8289 | -82.8916
5.8292 | -82.8895
5.8295 | -82.8873
5.8299 | -82.886
5.8302 | -82.8847
5.8305 | -82.8838
5.8308 | -82.8831
5.8311 | -82.8821
5.8315 | -82.8795
5.8318 | -82.8784
5.8321 | -82.8776
5.8324 | -82.8766
5.8328 | -82.8762
5.8331 | -82.8755
5.8334 | -82.8742
5.8337 | -82.8731
5.834 | -82.8723
5.8344 | -82.8713
5.8347 | -82.8712
5.835 | -82.8704
5.8353 | -82.8703
5.8356 | -82.8695
5.836 | -82.8685
5.8363 | -82.8675
5.8366 | -82.8669
5.8369 | -82.866
5.8373 | -82.8652
5.8376 | -82.8645
5.8379 | -82.8635
5.8382 | -82.8627
5.8385 | -82.8625
5.8389 | -82.8628
5.8392 | -82.8627
5.8395 | -82.8618
5.8398 | -82.861
5.8401 | -82.8604
5.8405 | -82.8594
5.8408 | -82.858
5.8411 | -82.8572
5.8414 | -82.8563
5.8418 | -82.8561
5.8421 | -82.8548
5.8424 | -82.8538
5.8427 | -82.853
5.843 | -82.8522
5.8434 | -82.851
5.8437 | -82.8496
5.844 | -82.8489
5.8443 | -82.8485
5.8446 | -82.8476
5.845 | -82.8469
5.8453 | -82.8466
5.8456 | -82.8456
5.8459 | -82.8445
5.8463 | -82.8439
5.8466 | -82.8435
5.8469 | -82.8433
5.8472 | -82.843
5.8475 | -82.8425
5.8479 | -82.8419
5.8482 | -82.8412
5.8485 | -82.8402
5.8488 | -82.8394
5.8491 | -82.8383
5.8495 | -82.8362
5.8498 | -82.8345
5.8501 | -82.8337
5.8504 | -82.8327
5.8507 | -82.8315
5.8511 | -82.8308
5.8514 | -82.8301
5.8517 | -82.829
5.852 | -82.8279
5.8524 | -82.826
5.8527 | -82.8252
5.853 | -82.8245
5.8533 | -82.8237
5.8536 | -82.8229
5.854 | -82.8226
5.8543 | -82.822
5.8546 | -82.8208
5.8549 | -82.82
5.8552 | -82.8194
5.8556 | -82.8183
5.8559 | -82.8169
5.8562 | -82.8164
5.8565 | -82.8158
5.8569 | -82.8151
5.8572 | -82.8146
5.8575 | -82.8138
5.8578 | -82.813
5.8581 | -82.8123
5.8585 | -82.8117
5.8588 | -82.8108
5.8591 | -82.8105
5.8594 | -82.8099
5.8597 | -82.8092
5.8601 | -82.8085
5.8604 | -82.8077
5.8607 | -82.8067
5.861 | -82.8052
5.8614 | -82.804
5.8617 | -82.804
5.862 | -82.8033
5.8623 | -82.8035
5.8626 | -82.8029
5.863 | -82.802
5.8633 | -82.8008
5.8636 | -82.8004
5.8639 | -82.8003
5.8642 | -82.7997
5.8646 | -82.7991
5.8649 | -82.7983
5.8652 | -82.798
5.8655 | -82.7975
5.8659 | -82.7967
5.8662 | -82.796
5.8665 | -82.7948
5.8668 | -82.7933
5.8671 | -82.7923
5.8675 | -82.7912
5.8678 | -82.7905
5.8681 | -82.79
5.8684 | -82.7901
5.8687 | -82.7898
5.8691 | -82.7902
5.8694 | -82.7903
5.8697 | -82.7904
5.87 | -82.7898
5.8704 | -82.7894
5.8707 | -82.7892
5.871 | -82.7883
5.8713 | -82.7875
5.8716 | -82.7873
5.872 | -82.7868
5.8723 | -82.7859
5.8726 | -82.7851
5.8729 | -82.7842
5.8732 | -82.7833
5.8736 | -82.7826
5.8739 | -82.7821
5.8742 | -82.7812
5.8745 | -82.7804
5.8748 | -82.7801
5.8752 | -82.7795
5.8755 | -82.7789
5.8758 | -82.7788
5.8761 | -82.7783
5.8765 | -82.7776
5.8768 | -82.7776
5.8771 | -82.778
5.8774 | -82.7777
5.8777 | -82.7779
5.8781 | -82.7776
5.8784 | -82.7776
5.8787 | -82.7782
5.879 | -82.7787
5.8793 | -82.7781
5.8797 | -82.7772
5.88 | -82.7762
5.8803 | -82.7755
5.8806 | -82.775
5.881 | -82.7746
5.8813 | -82.7745
5.8816 | -82.7748
5.8819 | -82.7745
5.8822 | -82.7737
5.8826 | -82.7728
5.8829 | -82.7714
5.8832 | -82.7706
5.8835 | -82.7704
5.8838 | -82.7698
5.8842 | -82.7693
5.8845 | -82.7693
5.8848 | -82.7698
5.8851 | -82.7693
5.8855 | -82.769
5.8858 | -82.7683
5.8861 | -82.7672
5.8864 | -82.7669
5.8867 | -82.766
5.8871 | -82.7659
5.8874 | -82.7656
5.8877 | -82.7656
5.888 | -82.7661
5.8883 | -82.7659
5.8887 | -82.7654
5.889 | -82.7651
5.8893 | -82.7645
5.8896 | -82.7645
5.89 | -82.764
5.8903 | -82.7641
5.8906 | -82.7643
5.8909 | -82.7642
5.8912 | -82.7639
5.8916 | -82.7627
5.8919 | -82.7616
5.8922 | -82.7614
5.8925 | -82.7605
5.8928 | -82.7606
5.8932 | -82.7614
5.8935 | -82.7615
5.8938 | -82.7619
5.8941 | -82.7618
5.8944 | -82.7616
5.8948 | -82.7616
5.8951 | -82.7607
5.8954 | -82.7597
5.8957 | -82.7596
5.8961 | -82.7598
5.8964 | -82.76
5.8967 | -82.7595
5.897 | -82.7594
5.8973 | -82.7594
5.8977 | -82.7601
5.898 | -82.7612
5.8983 | -82.7611
5.8986 | -82.7607
5.8989 | -82.7602
5.8993 | -82.7601
5.8996 | -82.7601
5.8999 | -82.76
5.9002 | -82.7599
5.9006 | -82.7598
5.9009 | -82.7599
5.9012 | -82.7597
5.9015 | -82.7596
5.9018 | -82.7592
5.9022 | -82.7593
5.9025 | -82.759
5.9028 | -82.7586
5.9031 | -82.7583
5.9034 | -82.7576
5.9038 | -82.757
5.9041 | -82.7573
5.9044 | -82.7571
5.9047 | -82.7568
5.9051 | -82.7568
5.9054 | -82.7567
5.9057 | -82.7574
5.906 | -82.7575
5.9063 | -82.7572
5.9067 | -82.7567
5.907 | -82.7569
5.9073 | -82.7569
5.9076 | -82.757
5.9079 | -82.7573
5.9083 | -82.7572
5.9086 | -82.7571
5.9089 | -82.7572
5.9092 | -82.7575
5.9096 | -82.7581
5.9099 | -82.7586
5.9102 | -82.7586
5.9105 | -82.7589
5.9108 | -82.7595
5.9112 | -82.7595
5.9115 | -82.7595
5.9118 | -82.7593
5.9121 | -82.7594
5.9124 | -82.7593
5.9128 | -82.76
5.9131 | -82.7607
5.9134 | -82.7607
5.9137 | -82.7607
5.9141 | -82.7606
5.9144 | -82.7603
5.9147 | -82.7606
5.915 | -82.7607
5.9153 | -82.7609
5.9157 | -82.7613
5.916 | -82.7609
5.9163 | -82.7604
5.9166 | -82.761
5.9169 | -82.7619
5.9173 | -82.7626
5.9176 | -82.7626
5.9179 | -82.7621
5.9182 | -82.7618
5.9185 | -82.7627
5.9189 | -82.7632
5.9192 | -82.7632
5.9195 | -82.7638
5.9198 | -82.7639
5.9202 | -82.7629
5.9205 | -82.7632
5.9208 | -82.763
5.9211 | -82.763
5.9214 | -82.7632
5.9218 | -82.763
5.9221 | -82.763
5.9224 | -82.7618
5.9227 | -82.7606
5.923 | -82.7612
5.9234 | -82.7613
5.9237 | -82.7614
5.924 | -82.7626
5.9243 | -82.763
5.9247 | -82.7636
5.925 | -82.7636
5.9253 | -82.764
5.9256 | -82.7647
5.9259 | -82.765
5.9263 | -82.7658
5.9266 | -82.7661
5.9269 | -82.7666
5.9272 | -82.7668
5.9275 | -82.7677
5.9279 | -82.7686
5.9282 | -82.7694
5.9285 | -82.7707
5.9288 | -82.7711
5.9292 | -82.7712
5.9295 | -82.7715
5.9298 | -82.7714
5.9301 | -82.7716
5.9304 | -82.7726
5.9308 | -82.7732
5.9311 | -82.7739
5.9314 | -82.7744
5.9317 | -82.7751
5.932 | -82.7758
5.9324 | -82.7761
5.9327 | -82.7763
5.933 | -82.7767
5.9333 | -82.7778
5.9337 | -82.7784
5.934 | -82.7793
5.9343 | -82.7797
5.9346 | -82.78
5.9349 | -82.7803
5.9353 | -82.7811
5.9356 | -82.7818
5.9359 | -82.7824
5.9362 | -82.7826
5.9365 | -82.783
5.9369 | -82.7834
5.9372 | -82.7831
5.9375 | -82.7839
5.9378 | -82.7843
5.9382 | -82.7851
5.9385 | -82.7855
5.9388 | -82.7854
5.9391 | -82.785
5.9394 | -82.7851
5.9398 | -82.7853
5.9401 | -82.7856
5.9404 | -82.7864
5.9407 | -82.7867
5.941 | -82.7873
5.9414 | -82.788
5.9417 | -82.7885
5.942 | -82.7888
5.9423 | -82.789
5.9426 | -82.7902
5.943 | -82.7909
5.9433 | -82.7922
5.9436 | -82.7929
5.9439 | -82.7932
5.9443 | -82.7935
5.9446 | -82.7945
5.9449 | -82.7955
5.9452 | -82.7969
5.9455 | -82.7976
5.9459 | -82.7983
5.9462 | -82.7992
5.9465 | -82.7997
5.9468 | -82.8002
5.9471 | -82.8012
5.9475 | -82.8019
5.9478 | -82.8026
5.9481 | -82.8037
5.9484 | -82.8038
5.9488 | -82.8034
5.9491 | -82.804
5.9494 | -82.8046
5.9497 | -82.8049
5.95 | -82.8052
5.9504 | -82.8062
5.9507 | -82.8073
5.951 | -82.8081
5.9513 | -82.8089
5.9516 | -82.8095
5.952 | -82.8104
5.9523 | -82.8112
5.9526 | -82.8119
5.9529 | -82.8125
5.9533 | -82.8134
5.9536 | -82.8139
5.9539 | -82.815
5.9542 | -82.8162
5.9545 | -82.8173
5.9549 | -82.8181
5.9552 | -82.8194
5.9555 | -82.8211
5.9558 | -82.8223
5.9561 | -82.823
5.9565 | -82.824
5.9568 | -82.8252
5.9571 | -82.8261
5.9574 | -82.8267
5.9578 | -82.8274
5.9581 | -82.8282
5.9584 | -82.8291
5.9587 | -82.8301
5.959 | -82.8309
5.9594 | -82.8314
5.9597 | -82.8327
5.96 | -82.8342
5.9603 | -82.835
5.9606 | -82.8362
5.961 | -82.8371
5.9613 | -82.8388
5.9616 | -82.8395
5.9619 | -82.8404
5.9623 | -82.8425
5.9626 | -82.8438
5.9629 | -82.8451
5.9632 | -82.8458
5.9635 | -82.8466
5.9639 | -82.8476
5.9642 | -82.8485
5.9645 | -82.8495
5.9648 | -82.8502
5.9651 | -82.8512
5.9655 | -82.852
5.9658 | -82.8538
5.9661 | -82.8555
5.9664 | -82.8561
5.9667 | -82.8568
5.9671 | -82.8579
5.9674 | -82.859
5.9677 | -82.8604
5.968 | -82.8624
5.9684 | -82.8648
5.9687 | -82.8662
5.969 | -82.8676
5.9693 | -82.8686
5.9696 | -82.8694
5.97 | -82.8697
5.9703 | -82.8709
5.9706 | -82.8717
5.9709 | -82.8728
5.9712 | -82.874
5.9716 | -82.8752
5.9719 | -82.8763
5.9722 | -82.8775
5.9725 | -82.8783
5.9729 | -82.8799
5.9732 | -82.8817
5.9735 | -82.8831
5.9738 | -82.8842
5.9741 | -82.8851
5.9745 | -82.8864
5.9748 | -82.888
5.9751 | -82.8892
5.9754 | -82.8903
5.9757 | -82.8916
5.9761 | -82.8928
5.9764 | -82.8935
5.9767 | -82.8942
5.977 | -82.895
5.9774 | -82.8952
5.9777 | -82.8963
5.978 | -82.8973
5.9783 | -82.898
5.9786 | -82.8992
5.979 | -82.9006
5.9793 | -82.902
5.9796 | -82.9043
5.9799 | -82.9064
5.9802 | -82.9082
5.9806 | -82.9097
5.9809 | -82.911
5.9812 | -82.9123
5.9815 | -82.9134
5.9819 | -82.9147
5.9822 | -82.916
5.9825 | -82.9171
5.9828 | -82.9183
5.9831 | -82.9196
5.9835 | -82.9209
5.9838 | -82.9223
5.9841 | -82.9244
5.9844 | -82.9256
5.9847 | -82.9269
5.9851 | -82.9283
5.9854 | -82.9298
5.9857 | -82.9311
5.986 | -82.9325
5.9863 | -82.9343
5.9867 | -82.9359
5.987 | -82.9374
5.9873 | -82.939
5.9876 | -82.9413
5.988 | -82.9432
5.9883 | -82.9439
5.9886 | -82.9456
5.9889 | -82.9481
5.9892 | -82.9497
5.9896 | -82.9514
5.9899 | -82.9542
5.9902 | -82.9572
5.9905 | -82.9591
5.9908 | -82.9602
5.9912 | -82.9613
5.9915 | -82.9619
5.9918 | -82.9638
5.9921 | -82.9663
5.9925 | -82.9684
5.9928 | -82.97
5.9931 | -82.9716
5.9934 | -82.9734
5.9937 | -82.975
5.9941 | -82.9768
5.9944 | -82.9784
5.9947 | -82.9798
5.995 | -82.9813
5.9953 | -82.9836
5.9957 | -82.9851
5.996 | -82.9868
5.9963 | -82.9884
5.9966 | -82.9893
5.997 | -82.9904
5.9973 | -82.9925
5.9976 | -82.9949
5.9979 | -82.9969
5.9982 | -82.9997
5.9986 | -83.0015
5.9989 | -83.0032
5.9992 | -83.0049
5.9995 | -83.0065
5.9998 | -83.008
6.0002 | -83.0093
6.0005 | -83.0105
6.0008 | -83.0121
6.0011 | -83.0147
6.0015 | -83.0168
6.0018 | -83.0178
6.0021 | -83.0192
6.0024 | -83.0211
6.0027 | -83.0233
6.0031 | -83.0256
6.0034 | -83.0278
6.0037 | -83.0301
6.004 | -83.0324
6.0043 | -83.0347
6.0047 | -83.0363
6.005 | -83.0378
6.0053 | -83.0406
6.0056 | -83.0425
6.006 | -83.0444
6.0063 | -83.0463
6.0066 | -83.0484
6.0069 | -83.0505
6.0072 | -83.0517
6.0076 | -83.0542
6.0079 | -83.0562
6.0082 | -83.0582
6.0085 | -83.0595
6.0088 | -83.0611
6.0092 | -83.063
6.0095 | -83.0648
6.0098 | -83.0668
6.0101 | -83.0689
6.0104 | -83.0709
6.0108 | -83.0729
6.0111 | -83.076
6.0114 | -83.0788
6.0117 | -83.0807
6.0121 | -83.0821
6.0124 | -83.0846
6.0127 | -83.0869
6.013 | -83.0889
6.0133 | -83.0911
6.0137 | -83.0928
6.014 | -83.0946
6.0143 | -83.0966
6.0146 | -83.0985
6.0149 | -83.1004
6.0153 | -83.1024
6.0156 | -83.1045
6.0159 | -83.1062
6.0162 | -83.1082
6.0166 | -83.1101
6.0169 | -83.1127
6.0172 | -83.1148
6.0175 | -83.1172
6.0178 | -83.1194
6.0182 | -83.1218
6.0185 | -83.1259
6.0188 | -83.1296
6.0191 | -83.132
6.0194 | -83.1339
6.0198 | -83.1369
6.0201 | -83.1394
6.0204 | -83.1414
6.0207 | -83.1434
6.0211 | -83.1453
6.0214 | -83.1469
6.0217 | -83.148
6.022 | -83.1495
6.0223 | -83.1513
6.0227 | -83.1533
6.023 | -83.1552
6.0233 | -83.1573
6.0236 | -83.1596
6.0239 | -83.1627
6.0243 | -83.1651
6.0246 | -83.1672
6.0249 | -83.1691
6.0252 | -83.1713
6.0256 | -83.1742
6.0259 | -83.1759
6.0262 | -83.1786
6.0265 | -83.1807
6.0268 | -83.1827
6.0272 | -83.1854
6.0275 | -83.1879
6.0278 | -83.19
6.0281 | -83.1927
6.0284 | -83.195
6.0288 | -83.1971
6.0291 | -83.1993
6.0294 | -83.202
6.0297 | -83.2045
6.0301 | -83.2068
6.0304 | -83.2087
6.0307 | -83.2105
6.031 | -83.2133
6.0313 | -83.216
6.0317 | -83.2181
6.032 | -83.2205
6.0323 | -83.223
6.0326 | -83.2261
6.0329 | -83.2285
6.0333 | -83.2307
6.0336 | -83.2329
6.0339 | -83.2357
6.0342 | -83.2384
6.0345 | -83.2409
6.0349 | -83.2435
6.0352 | -83.2465
6.0355 | -83.2484
6.0358 | -83.2504
6.0362 | -83.253
6.0365 | -83.2564
6.0368 | -83.2592
6.0371 | -83.2608
6.0374 | -83.2615
6.0378 | -83.263
6.0381 | -83.2653
6.0384 | -83.2677
6.0387 | -83.2699
6.039 | -83.2732
6.0394 | -83.2763
6.0397 | -83.2776
6.04 | -83.2799
6.0403 | -83.2821
6.0407 | -83.2846
6.041 | -83.287
6.0413 | -83.2895
6.0416 | -83.2918
6.0419 | -83.2937
6.0423 | -83.296
6.0426 | -83.2982
6.0429 | -83.2996
6.0432 | -83.3008
6.0435 | -83.303
6.0439 | -83.305
6.0442 | -83.3078
6.0445 | -83.3096
6.0448 | -83.3126
6.0452 | -83.3148
6.0455 | -83.3165
6.0458 | -83.3176
6.0461 | -83.3198
6.0464 | -83.3221
6.0468 | -83.3243
6.0471 | -83.327
6.0474 | -83.3294
6.0477 | -83.3317
6.048 | -83.3339
6.0484 | -83.3363
6.0487 | -83.3385
6.049 | -83.3409
6.0493 | -83.3434
6.0497 | -83.346
6.05 | -83.3486
6.0503 | -83.3509
6.0506 | -83.3533
6.0509 | -83.3556
6.0513 | -83.3573
6.0516 | -83.3598
6.0519 | -83.3629
6.0522 | -83.3653
6.0525 | -83.3672
6.0529 | -83.3686
6.0532 | -83.3705
6.0535 | -83.3729
6.0538 | -83.3753
6.0541 | -83.378
6.0545 | -83.3806
6.0548 | -83.3836
6.0551 | -83.386
6.0554 | -83.3876
6.0558 | -83.3892
6.0561 | -83.3906
6.0564 | -83.3923
6.0567 | -83.3947
6.057 | -83.3969
6.0574 | -83.4004
6.0577 | -83.4029
6.058 | -83.4058
6.0583 | -83.4082
6.0586 | -83.411
6.059 | -83.4138
6.0593 | -83.4161
6.0596 | -83.418
6.0599 | -83.4204
6.0603 | -83.4227
6.0606 | -83.4253
6.0609 | -83.4281
6.0612 | -83.4307
6.0615 | -83.4332
6.0619 | -83.4354
6.0622 | -83.4379
6.0625 | -83.4409
6.0628 | -83.4437
6.0631 | -83.4468
6.0635 | -83.4489
6.0638 | -83.4517
6.0641 | -83.4549
6.0644 | -83.4579
6.0648 | -83.4601
6.0651 | -83.4626
6.0654 | -83.4647
6.0657 | -83.4684
6.066 | -83.472
6.0664 | -83.475
6.0667 | -83.4769
6.067 | -83.4791
6.0673 | -83.4817
6.0676 | -83.4837
6.068 | -83.4856
6.0683 | -83.4879
6.0686 | -83.4901
6.0689 | -83.4926
6.0693 | -83.4947
6.0696 | -83.4969
6.0699 | -83.4984
6.0702 | -83.5008
6.0705 | -83.5033
6.0709 | -83.5062
6.0712 | -83.509
6.0715 | -83.5116
6.0718 | -83.5148
6.0721 | -83.5173
6.0725 | -83.52
6.0728 | -83.5225
6.0731 | -83.5251
6.0734 | -83.5278
6.0738 | -83.5301
6.0741 | -83.5326
6.0744 | -83.5357
6.0747 | -83.5376
6.075 | -83.5401
6.0754 | -83.5433
6.0757 | -83.5455
6.076 | -83.5478
6.0763 | -83.5507
6.0766 | -83.5536
6.077 | -83.556
6.0773 | -83.5591
6.0776 | -83.5619
6.0779 | -83.5644
6.0782 | -83.5663
6.0786 | -83.5682
6.0789 | -83.5707
6.0792 | -83.5732
6.0795 | -83.5754
6.0799 | -83.5778
6.0802 | -83.5802
6.0805 | -83.5823
6.0808 | -83.5842
6.0811 | -83.5863
6.0815 | -83.5887
6.0818 | -83.591
6.0821 | -83.5938
6.0824 | -83.5964
6.0827 | -83.5994
6.0831 | -83.602
6.0834 | -83.6052
6.0837 | -83.6076
6.084 | -83.6099
6.0844 | -83.6129
6.0847 | -83.6152
6.085 | -83.6174
6.0853 | -83.62
6.0856 | -83.6223
6.086 | -83.6247
6.0863 | -83.6274
6.0866 | -83.6288
6.0869 | -83.6304
6.0872 | -83.6338
6.0876 | -83.6378
6.0879 | -83.6438
6.0882 | -83.6479
6.0885 | -83.6505
6.0889 | -83.6535
6.0892 | -83.6579
6.0895 | -83.6615
6.0898 | -83.6629
6.0901 | -83.6649
6.0905 | -83.6668
6.0908 | -83.6693
6.0911 | -83.6717
6.0914 | -83.6744
6.0917 | -83.6764
6.0921 | -83.6788
6.0924 | -83.6815
6.0927 | -83.6838
6.093 | -83.6864
6.0934 | -83.6886
6.0937 | -83.6908
6.094 | -83.6918
6.0943 | -83.6916
6.0946 | -83.6914
6.095 | -83.6926
6.0953 | -83.6946
6.0956 | -83.6971
6.0959 | -83.6993
6.0962 | -83.7016
6.0966 | -83.7036
6.0969 | -83.7057
6.0972 | -83.7078
6.0975 | -83.7096
6.0979 | -83.7125
6.0982 | -83.7151
6.0985 | -83.718
6.0988 | -83.7211
6.0991 | -83.7235
6.0995 | -83.7258
6.0998 | -83.728
6.1001 | -83.7303
6.1004 | -83.7329
6.1007 | -83.7355
6.1011 | -83.7377
6.1014 | -83.7398
6.1017 | -83.7421
6.102 | -83.7443
6.1024 | -83.7458
6.1027 | -83.7483
6.103 | -83.7513
6.1033 | -83.7535
6.1036 | -83.756
6.104 | -83.7585
6.1043 | -83.7605
6.1046 | -83.7626
6.1049 | -83.7652
6.1052 | -83.7679
6.1056 | -83.7703
6.1059 | -83.772
6.1062 | -83.7729
6.1065 | -83.7736
6.1068 | -83.7745
6.1072 | -83.7767
6.1075 | -83.7785
6.1078 | -83.7809
6.1081 | -83.783
6.1085 | -83.7856
6.1088 | -83.7887
6.1091 | -83.7918
6.1094 | -83.7911
6.1097 | -83.7942
6.1101 | -83.7968
6.1104 | -83.7978
6.1107 | -83.7994
6.111 | -83.8011
6.1113 | -83.8029
6.1117 | -83.8052
6.112 | -83.807
6.1123 | -83.809
6.1126 | -83.811
6.113 | -83.8133
6.1133 | -83.8163
6.1136 | -83.8181
6.1139 | -83.8207
6.1142 | -83.8229
6.1146 | -83.8253
6.1149 | -83.8285
6.1152 | -83.832
6.1155 | -83.8343
6.1158 | -83.8377
6.1162 | -83.8412
6.1165 | -83.8435
6.1168 | -83.8458
6.1171 | -83.8479
6.1175 | -83.8497
6.1178 | -83.8519
6.1181 | -83.8541
6.1184 | -83.8579
6.1187 | -83.8595
6.1191 | -83.8626
6.1194 | -83.8648
6.1197 | -83.8663
6.12 | -83.8687
6.1203 | -83.8709
6.1207 | -83.8727
6.121 | -83.8746
6.1213 | -83.8763
6.1216 | -83.8786
6.122 | -83.8817
6.1223 | -83.8831
6.1226 | -83.8847
6.1229 | -83.8866
6.1232 | -83.8888
6.1236 | -83.8907
6.1239 | -83.8929
6.1242 | -83.8952
6.1245 | -83.8979
6.1248 | -83.8992
6.1252 | -83.9004
6.1255 | -83.9021
6.1258 | -83.9041
6.1261 | -83.9059
6.1264 | -83.9078
6.1268 | -83.9096
6.1271 | -83.9115
6.1274 | -83.9131
6.1277 | -83.9149
6.1281 | -83.9163
6.1284 | -83.9182
6.1287 | -83.9196
6.129 | -83.92
6.1293 | -83.9217
6.1297 | -83.9238
6.13 | -83.9256
6.1303 | -83.9276
6.1306 | -83.9289
6.1309 | -83.9301
6.1313 | -83.9311
6.1316 | -83.9324
6.1319 | -83.9344
6.1322 | -83.9358
6.1326 | -83.9368
6.1329 | -83.9383
6.1332 | -83.9414
6.1335 | -83.9471
6.1338 | -83.9507
6.1342 | -83.9525
6.1345 | -83.954
6.1348 | -83.956
6.1351 | -83.9597
6.1354 | -83.9621
6.1358 | -83.9648
6.1361 | -83.9665
6.1364 | -83.9674
6.1367 | -83.9696
6.1371 | -83.9711
6.1374 | -83.9722
6.1377 | -83.9745
6.138 | -83.9764
6.1383 | -83.9778
6.1387 | -83.9795
6.139 | -83.9817
6.1393 | -83.9836
6.1396 | -83.9856
6.1399 | -83.9876
6.1403 | -83.9887
6.1406 | -83.9904
6.1409 | -83.9919
6.1412 | -83.9918
6.1416 | -83.9913
6.1419 | -83.9931
6.1422 | -83.9954
6.1425 | -83.9964
6.1428 | -83.9986
6.1432 | -84.0002
6.1435 | -84.0007
6.1438 | -84.0011
6.1441 | -84.0028
6.1444 | -84.0045
6.1448 | -84.0051
6.1451 | -84.0059
6.1454 | -84.0073
6.1457 | -84.0086
6.146 | -84.0097
6.1464 | -84.01
6.1467 | -84.0094
6.147 | -84.0115
6.1473 | -84.0127
6.1477 | -84.0142
6.148 | -84.015
6.1483 | -84.0159
6.1486 | -84.0172
6.1489 | -84.0184
6.1493 | -84.0197
6.1496 | -84.0204
6.1499 | -84.0202
6.1502 | -84.0199
6.1505 | -84.021
6.1509 | -84.0219
6.1512 | -84.0226
6.1515 | -84.024
6.1518 | -84.0254
6.1522 | -84.0262
6.1525 | -84.0271
6.1528 | -84.0279
6.1531 | -84.0293
6.1534 | -84.0307
6.1538 | -84.0317
6.1541 | -84.0325
6.1544 | -84.034
6.1547 | -84.0347
6.155 | -84.0347
6.1554 | -84.0348
6.1557 | -84.0358
6.156 | -84.0364
6.1563 | -84.0375
6.1567 | -84.0387
6.157 | -84.0406
6.1573 | -84.0417
6.1576 | -84.043
6.1579 | -84.0452
6.1583 | -84.0463
6.1586 | -84.0469
6.1589 | -84.0459
6.1592 | -84.0453
6.1595 | -84.0459
6.1599 | -84.0466
6.1602 | -84.0471
6.1605 | -84.0472
6.1608 | -84.0491
6.1612 | -84.0501
6.1615 | -84.0503
6.1618 | -84.0509
6.1621 | -84.0516
6.1624 | -84.0529
6.1628 | -84.056
6.1631 | -84.057
6.1634 | -84.0579
6.1637 | -84.0603
6.164 | -84.0612
6.1644 | -84.0614
6.1647 | -84.062
6.165 | -84.0617
6.1653 | -84.0618
6.1657 | -84.0628
6.166 | -84.0646
6.1663 | -84.0657
6.1666 | -84.067
6.1669 | -84.0679
6.1673 | -84.0679
6.1676 | -84.0675
6.1679 | -84.0677
6.1682 | -84.0679
6.1685 | -84.0686
6.1689 | -84.07
6.1692 | -84.0711
6.1695 | -84.072
6.1698 | -84.072
6.1701 | -84.072
6.1705 | -84.0726
6.1708 | -84.0735
6.1711 | -84.0739
6.1714 | -84.0741
6.1718 | -84.0745
6.1721 | -84.0753
6.1724 | -84.076
6.1727 | -84.0764
6.173 | -84.0769
6.1734 | -84.0775
6.1737 | -84.0782
6.174 | -84.0786
6.1743 | -84.0789
6.1746 | -84.0793
6.175 | -84.0801
6.1753 | -84.0811
6.1756 | -84.0819
6.1759 | -84.0827
6.1763 | -84.084
6.1766 | -84.0852
6.1769 | -84.0852
6.1772 | -84.0856
6.1775 | -84.0862
6.1779 | -84.0872
6.1782 | -84.0889
6.1785 | -84.091
6.1788 | -84.0935
6.1791 | -84.0953
6.1795 | -84.0968
6.1798 | -84.0969
6.1801 | -84.0977
6.1804 | -84.0979
6.1808 | -84.0977
6.1811 | -84.0977
6.1814 | -84.0979
6.1817 | -84.0984
6.182 | -84.0987
6.1824 | -84.099
6.1827 | -84.0994
6.183 | -84.0999
6.1833 | -84.1001
6.1836 | -84.0996
6.184 | -84.0985
6.1843 | -84.0984
6.1846 | -84.0987
6.1849 | -84.0982
6.1853 | -84.0977
6.1856 | -84.0976
6.1859 | -84.0974
6.1862 | -84.0976
6.1865 | -84.0977
6.1869 | -84.0981
6.1872 | -84.0983
6.1875 | -84.0982
6.1878 | -84.0992
6.1881 | -84.1021
6.1885 | -84.1027
6.1888 | -84.1013
6.1891 | -84.1016
6.1894 | -84.1018
6.1898 | -84.1015
6.1901 | -84.1021
6.1904 | -84.1029
6.1907 | -84.103
6.191 | -84.1031
6.1914 | -84.1033
6.1917 | -84.1036
6.192 | -84.1037
6.1923 | -84.1037
6.1926 | -84.1035
6.193 | -84.1017
6.1933 | -84.1017
6.1936 | -84.1023
6.1939 | -84.1033
6.1943 | -84.1046
6.1946 | -84.1067
6.1949 | -84.1083
6.1952 | -84.1077
6.1955 | -84.1076
6.1959 | -84.1086
6.1962 | -84.1091
6.1965 | -84.1091
6.1968 | -84.1103
6.1971 | -84.1124
6.1975 | -84.1148
6.1978 | -84.1154
6.1981 | -84.1153
6.1984 | -84.116
6.1987 | -84.1158
6.1991 | -84.116
6.1994 | -84.1158
6.1997 | -84.1163
6.2 | -84.1165
6.2004 | -84.1168
6.2007 | -84.1169
6.201 | -84.1171
6.2013 | -84.1173
6.2016 | -84.1168
6.202 | -84.117
6.2023 | -84.1168
6.2026 | -84.1169
6.2029 | -84.117
6.2032 | -84.1171
6.2036 | -84.1172
6.2039 | -84.1178
6.2042 | -84.118
6.2045 | -84.1177
6.2049 | -84.118
6.2052 | -84.1174
6.2055 | -84.1162
6.2058 | -84.117
6.2061 | -84.1173
6.2065 | -84.1179
6.2068 | -84.1184
6.2071 | -84.1188
6.2074 | -84.1187
6.2077 | -84.118
6.2081 | -84.1171
6.2084 | -84.1161
6.2087 | -84.1165
6.209 | -84.1164
6.2094 | -84.1167
6.2097 | -84.1172
6.21 | -84.1169
6.2103 | -84.1166
6.2106 | -84.1172
6.211 | -84.1175
6.2113 | -84.117
6.2116 | -84.1167
6.2119 | -84.116
6.2122 | -84.1153
6.2126 | -84.1151
6.2129 | -84.1149
6.2132 | -84.1152
6.2135 | -84.1157
6.2139 | -84.1156
6.2142 | -84.1152
6.2145 | -84.1145
6.2148 | -84.1143
6.2151 | -84.1153
6.2155 | -84.1161
6.2158 | -84.1163
6.2161 | -84.1159
6.2164 | -84.1158
6.2167 | -84.1154
6.2171 | -84.1151
6.2174 | -84.1151
6.2177 | -84.1151
6.218 | -84.1144
6.2183 | -84.114
6.2187 | -84.1146
6.219 | -84.1148
6.2193 | -84.1146
6.2196 | -84.1152
6.22 | -84.1146
6.2203 | -84.1146
6.2206 | -84.1145
6.2209 | -84.1146
6.2212 | -84.1143
6.2216 | -84.1134
6.2219 | -84.1129
6.2222 | -84.1132
6.2225 | -84.1133
6.2228 | -84.114
6.2232 | -84.1144
6.2235 | -84.1144
6.2238 | -84.1135
6.2241 | -84.1125
6.2245 | -84.112
6.2248 | -84.1113
6.2251 | -84.1109
6.2254 | -84.1108
6.2257 | -84.1112
6.2261 | -84.1106
6.2264 | -84.1106
6.2267 | -84.1101
6.227 | -84.1098
6.2273 | -84.1096
6.2277 | -84.1109
6.228 | -84.111
6.2283 | -84.1115
6.2286 | -84.1112
6.229 | -84.1112
6.2293 | -84.112
6.2296 | -84.1118
6.2299 | -84.1109
6.2302 | -84.1106
6.2306 | -84.1101
6.2309 | -84.1095
6.2312 | -84.1092
6.2315 | -84.1101
6.2318 | -84.1098
6.2322 | -84.1093
6.2325 | -84.11
6.2328 | -84.1109
6.2331 | -84.1104
6.2335 | -84.1105
6.2338 | -84.111
6.2341 | -84.1108
6.2344 | -84.1106
6.2347 | -84.1102
6.2351 | -84.111
6.2354 | -84.1105
6.2357 | -84.1095
6.236 | -84.1089
6.2363 | -84.1083
6.2367 | -84.1077
6.237 | -84.107
6.2373 | -84.1075
6.2376 | -84.1084
6.2379 | -84.1082
6.2383 | -84.1086
6.2386 | -84.109
6.2389 | -84.1087
6.2392 | -84.1084
6.2396 | -84.1077
6.2399 | -84.1066
6.2402 | -84.1059
6.2405 | -84.1051
6.2408 | -84.1047
6.2412 | -84.1039
6.2415 | -84.1035
6.2418 | -84.103
6.2421 | -84.1024
6.2424 | -84.1018
6.2428 | -84.1015
6.2431 | -84.1014
6.2434 | -84.1011
6.2437 | -84.1011
6.2441 | -84.1007
6.2444 | -84.0997
6.2447 | -84.0995
6.245 | -84.0987
6.2453 | -84.0978
6.2457 | -84.097
6.246 | -84.0961
6.2463 | -84.0955
6.2466 | -84.0946
6.2469 | -84.0935
6.2473 | -84.093
6.2476 | -84.0925
6.2479 | -84.0927
6.2482 | -84.0941
6.2486 | -84.0941
6.2489 | -84.0946
6.2492 | -84.0943
6.2495 | -84.0945
6.2498 | -84.0939
6.2502 | -84.0919
6.2505 | -84.0911
6.2508 | -84.0922
6.2511 | -84.0931
6.2514 | -84.0929
6.2518 | -84.092
6.2521 | -84.0908
6.2524 | -84.0891
6.2527 | -84.0881
6.2531 | -84.0882
6.2534 | -84.088
6.2537 | -84.0867
6.254 | -84.0851
6.2543 | -84.0837
6.2547 | -84.0829
6.255 | -84.0818
6.2553 | -84.0812
6.2556 | -84.0804
6.2559 | -84.0791
6.2563 | -84.0778
6.2566 | -84.0775
6.2569 | -84.0775
6.2572 | -84.077
6.2576 | -84.0765
6.2579 | -84.0768
6.2582 | -84.0766
6.2585 | -84.0755
6.2588 | -84.0747
6.2592 | -84.0743
6.2595 | -84.0739
6.2598 | -84.0734
6.2601 | -84.0725
6.2604 | -84.0721
6.2608 | -84.0724
6.2611 | -84.0728
6.2614 | -84.0721
6.2617 | -84.0718
6.2621 | -84.0713
6.2624 | -84.0713
6.2627 | -84.071
6.263 | -84.0702
6.2633 | -84.07
6.2637 | -84.0699
6.264 | -84.0705
6.2643 | -84.0703
6.2646 | -84.0699
6.2649 | -84.0695
6.2653 | -84.0689
6.2656 | -84.0674
6.2659 | -84.066
6.2662 | -84.0658
6.2665 | -84.0654
6.2669 | -84.065
6.2672 | -84.0654
6.2675 | -84.0648
6.2678 | -84.0644
6.2682 | -84.064
6.2685 | -84.0638
6.2688 | -84.0636
6.2691 | -84.0635
6.2694 | -84.063
6.2698 | -84.0632
6.2701 | -84.0645
6.2704 | -84.0652
6.2707 | -84.0656
6.271 | -84.066
6.2714 | -84.0651
6.2717 | -84.0645
6.272 | -84.0642
6.2723 | -84.0643
6.2727 | -84.0643
6.273 | -84.0644
6.2733 | -84.0653
6.2736 | -84.0644
6.2739 | -84.0629
6.2743 | -84.0621
6.2746 | -84.0618
6.2749 | -84.0616
6.2752 | -84.0617
6.2755 | -84.0618
6.2759 | -84.0612
6.2762 | -84.061
6.2765 | -84.0608
6.2768 | -84.0611
6.2772 | -84.0612
6.2775 | -84.061
6.2778 | -84.0608
6.2781 | -84.0605
6.2784 | -84.0599
6.2788 | -84.0596
6.2791 | -84.0595
6.2794 | -84.0592
6.2797 | -84.0584
6.28 | -84.0577
6.2804 | -84.0566
6.2807 | -84.0559
6.281 | -84.0553
6.2813 | -84.0548
6.2817 | -84.0543
6.282 | -84.0541
6.2823 | -84.0542
6.2826 | -84.0539
6.2829 | -84.0547
6.2833 | -84.0546
6.2836 | -84.0545
6.2839 | -84.0542
6.2842 | -84.0543
6.2845 | -84.0542
6.2849 | -84.0543
6.2852 | -84.054
6.2855 | -84.0546
6.2858 | -84.0541
6.2861 | -84.0538
6.2865 | -84.0536
6.2868 | -84.0538
6.2871 | -84.0539
6.2874 | -84.0535
6.2878 | -84.0535
6.2881 | -84.0536
6.2884 | -84.0543
6.2887 | -84.0555
6.289 | -84.0559
6.2894 | -84.0561
6.2897 | -84.0557
6.29 | -84.0546
6.2903 | -84.0534
6.2906 | -84.0529
6.291 | -84.0522
6.2913 | -84.0518
6.2916 | -84.0515
6.2919 | -84.0512
6.2923 | -84.0513
6.2926 | -84.0511
6.2929 | -84.0509
6.2932 | -84.051
6.2935 | -84.0507
6.2939 | -84.0493
6.2942 | -84.0475
6.2945 | -84.0475
6.2948 | -84.0476
6.2951 | -84.0477
6.2955 | -84.048
6.2958 | -84.0479
6.2961 | -84.0467
6.2964 | -84.0464
6.2968 | -84.0465
6.2971 | -84.0466
6.2974 | -84.0458
6.2977 | -84.046
6.298 | -84.046
6.2984 | -84.0449
6.2987 | -84.0441
6.299 | -84.0431
6.2993 | -84.0425
6.2996 | -84.0428
6.3 | -84.0429
6.3003 | -84.0424
6.3006 | -84.0424
6.3009 | -84.0423
6.3013 | -84.0418
6.3016 | -84.0413
6.3019 | -84.0411
6.3022 | -84.0412
6.3025 | -84.0405
6.3029 | -84.04
6.3032 | -84.0398
6.3035 | -84.0397
6.3038 | -84.0393
6.3041 | -84.04
6.3045 | -84.0391
6.3048 | -84.0387
6.3051 | -84.0386
6.3054 | -84.0389
6.3057 | -84.0386
6.3061 | -84.0397
6.3064 | -84.0402
6.3067 | -84.0398
6.307 | -84.0394
6.3074 | -84.0395
6.3077 | -84.0399
6.308 | -84.041
6.3083 | -84.0419
6.3086 | -84.0421
6.309 | -84.0425
6.3093 | -84.0434
6.3096 | -84.0431
6.3099 | -84.0439
6.3102 | -84.0443
6.3106 | -84.0442
6.3109 | -84.0443
6.3112 | -84.0439
6.3115 | -84.0429
6.3119 | -84.0423
6.3122 | -84.0424
6.3125 | -84.0423
6.3128 | -84.0424
6.3131 | -84.0422
6.3135 | -84.0433
6.3138 | -84.0435
6.3141 | -84.0436
6.3144 | -84.0435
6.3147 | -84.0435
6.3151 | -84.0436
6.3154 | -84.0447
6.3157 | -84.0457
6.316 | -84.046
6.3164 | -84.0468
6.3167 | -84.047
6.317 | -84.0467
6.3173 | -84.0476
6.3176 | -84.0479
6.318 | -84.0476
6.3183 | -84.0468
6.3186 | -84.047
6.3189 | -84.0467
6.3192 | -84.0469
6.3196 | -84.0469
6.3199 | -84.0465
6.3202 | -84.0456
6.3205 | -84.0453
6.3209 | -84.0453
6.3212 | -84.0442
6.3215 | -84.0432
6.3218 | -84.0429
6.3221 | -84.0428
6.3225 | -84.0426
6.3228 | -84.043
6.3231 | -84.0428
6.3234 | -84.043
6.3237 | -84.0424
6.3241 | -84.0421
6.3244 | -84.042
6.3247 | -84.0421
6.325 | -84.0421
6.3254 | -84.0418
6.3257 | -84.0419
6.326 | -84.0426
6.3263 | -84.043
6.3266 | -84.0427
6.327 | -84.043
6.3273 | -84.0434
6.3276 | -84.0438
6.3279 | -84.0445
6.3282 | -84.0446
6.3286 | -84.0451
6.3289 | -84.0453
6.3292 | -84.0437
6.3295 | -84.042
6.3299 | -84.0422
6.3302 | -84.0427
6.3305 | -84.0423
6.3308 | -84.0423
6.3311 | -84.0425
6.3315 | -84.0421
6.3318 | -84.0417
6.3321 | -84.0419
6.3324 | -84.0416
6.3327 | -84.0418
6.3331 | -84.0411
6.3334 | -84.0398
6.3337 | -84.0393
6.334 | -84.0387
6.3343 | -84.0386
6.3347 | -84.0386
6.335 | -84.0385
6.3353 | -84.0375
6.3356 | -84.0373
6.336 | -84.0375
6.3363 | -84.0379
6.3366 | -84.0381
6.3369 | -84.0386
6.3372 | -84.0387
6.3376 | -84.0389
6.3379 | -84.0393
6.3382 | -84.0394
6.3385 | -84.0396
6.3388 | -84.0398
6.3392 | -84.0403
6.3395 | -84.041
6.3398 | -84.0417
6.3401 | -84.0415
6.3405 | -84.0412
6.3408 | -84.0412
6.3411 | -84.0405
6.3414 | -84.0405
6.3417 | -84.0414
6.3421 | -84.0422
6.3424 | -84.0424
6.3427 | -84.0427
6.343 | -84.0429
6.3433 | -84.0436
6.3437 | -84.0443
6.344 | -84.0449
6.3443 | -84.0449
6.3446 | -84.0447
6.345 | -84.0451
6.3453 | -84.0455
6.3456 | -84.0456
6.3459 | -84.0468
6.3462 | -84.0484
6.3466 | -84.0491
6.3469 | -84.0496
6.3472 | -84.0501
6.3475 | -84.0508
6.3478 | -84.051
6.3482 | -84.0523
6.3485 | -84.0543
6.3488 | -84.0549
6.3491 | -84.0552
6.3495 | -84.0555
6.3498 | -84.0559
6.3501 | -84.056
6.3504 | -84.0559
6.3507 | -84.0555
6.3511 | -84.0554
6.3514 | -84.0555
6.3517 | -84.0556
6.352 | -84.0559
6.3523 | -84.0567
6.3527 | -84.0571
6.353 | -84.0571
6.3533 | -84.0573
6.3536 | -84.0574
6.354 | -84.0578
6.3543 | -84.0582
6.3546 | -84.0585
6.3549 | -84.0587
6.3552 | -84.0588
6.3556 | -84.0588
6.3559 | -84.0586
6.3562 | -84.0595
6.3565 | -84.061
6.3568 | -84.0616
6.3572 | -84.0617
6.3575 | -84.0618
6.3578 | -84.062
6.3581 | -84.0619
6.3584 | -84.0625
6.3588 | -84.0632
6.3591 | -84.0639
6.3594 | -84.0642
6.3597 | -84.0652
6.3601 | -84.0656
6.3604 | -84.0664
6.3607 | -84.0671
6.361 | -84.0677
6.3613 | -84.0681
6.3617 | -84.0683
6.362 | -84.0689
6.3623 | -84.0695
6.3626 | -84.0702
6.3629 | -84.0709
6.3633 | -84.0713
6.3636 | -84.0721
6.3639 | -84.0733
6.3642 | -84.0742
6.3646 | -84.0747
6.3649 | -84.0754
6.3652 | -84.0757
6.3655 | -84.0764
6.3658 | -84.0777
6.3662 | -84.0789
6.3665 | -84.0799
6.3668 | -84.0806
6.3671 | -84.0817
6.3674 | -84.0823
6.3678 | -84.0827
6.3681 | -84.0827
6.3684 | -84.0832
6.3687 | -84.0835
6.3691 | -84.0847
6.3694 | -84.0854
6.3697 | -84.0861
6.37 | -84.0858
6.3703 | -84.086
6.3707 | -84.0866
6.371 | -84.087
6.3713 | -84.0871
6.3716 | -84.0878
6.3719 | -84.0885
6.3723 | -84.0885
6.3726 | -84.0886
6.3729 | -84.0887
6.3732 | -84.0889
6.3736 | -84.0888
6.3739 | -84.0884
6.3742 | -84.0887
6.3745 | -84.0894
6.3748 | -84.0899
6.3752 | -84.0899
6.3755 | -84.0904
6.3758 | -84.0908
6.3761 | -84.0909
6.3764 | -84.092
6.3768 | -84.0923
6.3771 | -84.0923
6.3774 | -84.0923
6.3777 | -84.0923
6.378 | -84.0917
6.3784 | -84.0914
6.3787 | -84.0908
6.379 | -84.0909
6.3793 | -84.0913
6.3797 | -84.0917
6.38 | -84.0925
6.3803 | -84.0932
6.3806 | -84.0932
6.3809 | -84.0931
6.3813 | -84.0932
6.3816 | -84.0932
6.3819 | -84.0929
6.3822 | -84.093
6.3825 | -84.0936
6.3829 | -84.0937
6.3832 | -84.0941
6.3835 | -84.0943
6.3838 | -84.0946
6.3842 | -84.0948
6.3845 | -84.0952
6.3848 | -84.0948
6.3851 | -84.0944
6.3854 | -84.0942
6.3858 | -84.0942
6.3861 | -84.0943
6.3864 | -84.0935
6.3867 | -84.0923
6.387 | -84.0925
6.3874 | -84.0935
6.3877 | -84.0942
6.388 | -84.0945
6.3883 | -84.0936
6.3887 | -84.0931
6.389 | -84.0933
6.3893 | -84.0939
6.3896 | -84.0943
6.3899 | -84.0943
6.3903 | -84.0947
6.3906 | -84.0951
6.3909 | -84.0955
6.3912 | -84.0943
6.3915 | -84.0933
6.3919 | -84.0928
6.3922 | -84.0936
6.3925 | -84.0941
6.3928 | -84.0944
6.3932 | -84.0947
6.3935 | -84.095
6.3938 | -84.0956
6.3941 | -84.0958
6.3944 | -84.0958
6.3948 | -84.0945
6.3951 | -84.0938
6.3954 | -84.0944
6.3957 | -84.0935
6.396 | -84.0914
6.3964 | -84.0905
6.3967 | -84.0905
6.397 | -84.0909
6.3973 | -84.0908
6.3976 | -84.0906
6.398 | -84.0913
6.3983 | -84.0908
6.3986 | -84.0904
6.3989 | -84.0909
6.3993 | -84.0921
6.3996 | -84.0931
6.3999 | -84.0939
6.4002 | -84.0943
6.4005 | -84.0943
6.4009 | -84.0943
6.4012 | -84.0944
6.4015 | -84.0947
6.4018 | -84.0945
6.4021 | -84.0943
6.4025 | -84.0943
6.4028 | -84.0937
6.4031 | -84.0934
6.4034 | -84.0931
6.4038 | -84.093
6.4041 | -84.0933
6.4044 | -84.0935
6.4047 | -84.0939
6.405 | -84.0942
6.4054 | -84.0944
6.4057 | -84.0949
6.406 | -84.0953
6.4063 | -84.0954
6.4066 | -84.0957
6.407 | -84.0962
6.4073 | -84.0975
6.4076 | -84.0985
6.4079 | -84.0992
6.4083 | -84.0997
6.4086 | -84.1
6.4089 | -84.0999
6.4092 | -84.0995
6.4095 | -84.0982
6.4099 | -84.0984
6.4102 | -84.0982
6.4105 | -84.0975
6.4108 | -84.097
6.4111 | -84.0963
6.4115 | -84.0962
6.4118 | -84.0967
6.4121 | -84.0968
6.4124 | -84.0966
6.4128 | -84.0969
6.4131 | -84.0973
6.4134 | -84.0976
6.4137 | -84.0976
6.414 | -84.0979
6.4144 | -84.098
6.4147 | -84.0976
6.415 | -84.0977
6.4153 | -84.0966
6.4156 | -84.096
6.416 | -84.0958
6.4163 | -84.0959
6.4166 | -84.0959
6.4169 | -84.0959
6.4173 | -84.0958
6.4176 | -84.0958
6.4179 | -84.0958
6.4182 | -84.0959
6.4185 | -84.0959
6.4189 | -84.0964
6.4192 | -84.0963
6.4195 | -84.0965
6.4198 | -84.0962
6.4201 | -84.0962
6.4205 | -84.0965
6.4208 | -84.0966
6.4211 | -84.0963
6.4214 | -84.0958
6.4218 | -84.0961
6.4221 | -84.0967
6.4224 | -84.0978
6.4227 | -84.098
6.423 | -84.0977
6.4234 | -84.0975
6.4237 | -84.0975
6.424 | -84.0976
6.4243 | -84.0975
6.4246 | -84.097
6.425 | -84.0962
6.4253 | -84.0958
6.4256 | -84.0956
6.4259 | -84.0958
6.4262 | -84.0958
6.4266 | -84.0956
6.4269 | -84.0959
6.4272 | -84.0958
6.4275 | -84.0959
6.4279 | -84.0959
6.4282 | -84.0963
6.4285 | -84.096
6.4288 | -84.0961
6.4291 | -84.0958
6.4295 | -84.0956
6.4298 | -84.0957
6.4301 | -84.0959
6.4304 | -84.0956
6.4307 | -84.0952
6.4311 | -84.0956
6.4314 | -84.0949
6.4317 | -84.0937
6.432 | -84.0943
6.4324 | -84.0954
6.4327 | -84.0951
6.433 | -84.0954
6.4333 | -84.0963
6.4336 | -84.0964
6.434 | -84.0961
6.4343 | -84.096
6.4346 | -84.0961
6.4349 | -84.0959
6.4352 | -84.0958
6.4356 | -84.0959
6.4359 | -84.0959
6.4362 | -84.0956
6.4365 | -84.0957
6.4369 | -84.0962
6.4372 | -84.096
6.4375 | -84.0954
6.4378 | -84.094
6.4381 | -84.0923
6.4385 | -84.092
6.4388 | -84.0914
6.4391 | -84.0917
6.4394 | -84.0918
6.4397 | -84.0915
6.4401 | -84.0905
6.4404 | -84.0896
6.4407 | -84.0901
6.441 | -84.0902
6.4414 | -84.09
6.4417 | -84.0896
6.442 | -84.089
6.4423 | -84.0889
6.4426 | -84.088
6.443 | -84.0873
6.4433 | -84.0866
6.4436 | -84.0858
6.4439 | -84.0857
6.4442 | -84.0858
6.4446 | -84.0865
6.4449 | -84.0866
6.4452 | -84.0863
6.4455 | -84.0859
6.4459 | -84.0857
6.4462 | -84.0856
6.4465 | -84.085
6.4468 | -84.0849
6.4471 | -84.0852
6.4475 | -84.0848
6.4478 | -84.0845
6.4481 | -84.0842
6.4484 | -84.0842
6.4487 | -84.0838
6.4491 | -84.0835
6.4494 | -84.0832
6.4497 | -84.0835
6.45 | -84.0827
6.4503 | -84.0815
6.4507 | -84.0813
6.451 | -84.0818
6.4513 | -84.0817
6.4516 | -84.0813
6.452 | -84.0809
6.4523 | -84.0808
6.4526 | -84.0797
6.4529 | -84.0787
6.4532 | -84.0789
6.4536 | -84.078
6.4539 | -84.077
6.4542 | -84.0769
6.4545 | -84.0773
6.4548 | -84.078
6.4552 | -84.0784
6.4555 | -84.0779
6.4558 | -84.0775
6.4561 | -84.0769
6.4565 | -84.0762
6.4568 | -84.0758
6.4571 | -84.0754
6.4574 | -84.0752
6.4577 | -84.075
6.4581 | -84.0748
6.4584 | -84.0745
6.4587 | -84.0745
6.459 | -84.0756
6.4593 | -84.0756
6.4597 | -84.0752
6.46 | -84.0748
6.4603 | -84.074
6.4606 | -84.0728
6.461 | -84.0717
6.4613 | -84.0716
6.4616 | -84.0712
6.4619 | -84.0698
6.4622 | -84.0685
6.4626 | -84.068
6.4629 | -84.0672
6.4632 | -84.0666
6.4635 | -84.0662
6.4638 | -84.0656
6.4642 | -84.0657
6.4645 | -84.0658
6.4648 | -84.0658
6.4651 | -84.0655
6.4655 | -84.065
6.4658 | -84.0643
6.4661 | -84.0638
6.4664 | -84.0633
6.4667 | -84.0628
6.4671 | -84.0621
6.4674 | -84.0612
6.4677 | -84.0604
6.468 | -84.0597
6.4683 | -84.0593
6.4687 | -84.0588
6.469 | -84.0585
6.4693 | -84.0573
6.4696 | -84.0571
6.4699 | -84.0568
6.4703 | -84.0564
6.4706 | -84.0557
6.4709 | -84.0553
6.4712 | -84.055
6.4716 | -84.0553
6.4719 | -84.0569
6.4722 | -84.0575
6.4725 | -84.0572
6.4728 | -84.0568
6.4732 | -84.056
6.4735 | -84.0546
6.4738 | -84.0538
6.4741 | -84.0544
6.4744 | -84.0532
6.4748 | -84.0518
6.4751 | -84.0507
6.4754 | -84.0501
6.4757 | -84.0496
6.4761 | -84.0494
6.4764 | -84.0491
6.4767 | -84.0484
6.477 | -84.0483
6.4773 | -84.0486
6.4777 | -84.0486
6.478 | -84.0478
6.4783 | -84.0473
6.4786 | -84.047
6.4789 | -84.0471
6.4793 | -84.0466
6.4796 | -84.0448
6.4799 | -84.0437
6.4802 | -84.0434
6.4806 | -84.0435
6.4809 | -84.0446
6.4812 | -84.0446
6.4815 | -84.0435
6.4818 | -84.0428
6.4822 | -84.0423
6.4825 | -84.0419
6.4828 | -84.0414
6.4831 | -84.0406
6.4834 | -84.0399
6.4838 | -84.0393
6.4841 | -84.0387
6.4844 | -84.0372
6.4847 | -84.0359
6.4851 | -84.0354
6.4854 | -84.0347
6.4857 | -84.0339
6.486 | -84.0334
6.4863 | -84.0336
6.4867 | -84.0329
6.487 | -84.0321
6.4873 | -84.0313
6.4876 | -84.0311
6.4879 | -84.0313
6.4883 | -84.0307
6.4886 | -84.0302
6.4889 | -84.0298
6.4892 | -84.0294
6.4896 | -84.0285
6.4899 | -84.0274
6.4902 | -84.0267
6.4905 | -84.0259
6.4908 | -84.0244
6.4912 | -84.0235
6.4915 | -84.0238
6.4918 | -84.0242
6.4921 | -84.0241
6.4924 | -84.0234
6.4928 | -84.0232
6.4931 | -84.0226
6.4934 | -84.0216
6.4937 | -84.0208
6.494 | -84.0199
6.4944 | -84.0187
6.4947 | -84.0182
6.495 | -84.0177
6.4953 | -84.0172
6.4957 | -84.0166
6.496 | -84.0158
6.4963 | -84.0146
6.4966 | -84.0135
6.4969 | -84.0128
6.4973 | -84.0116
6.4976 | -84.0104
6.4979 | -84.0096
6.4982 | -84.0087
6.4985 | -84.0071
6.4989 | -84.0063
6.4992 | -84.005
6.4995 | -84.0035
6.4998 | -84.0028
6.5002 | -84.0028
6.5005 | -84.002
6.5008 | -84.001
6.5011 | -83.9999
6.5014 | -83.9991
6.5018 | -83.9986
6.5021 | -83.9981
6.5024 | -83.9982
6.5027 | -83.9975
6.503 | -83.997
6.5034 | -83.9964
6.5037 | -83.9956
6.504 | -83.9954
6.5043 | -83.9955
6.5047 | -83.9949
6.505 | -83.9941
6.5053 | -83.993
6.5056 | -83.9914
6.5059 | -83.9913
6.5063 | -83.9914
6.5066 | -83.9907
6.5069 | -83.9899
6.5072 | -83.989
6.5075 | -83.9883
6.5079 | -83.9871
6.5082 | -83.9866
6.5085 | -83.9856
6.5088 | -83.9842
6.5092 | -83.9831
6.5095 | -83.9823
6.5098 | -83.9815
6.5101 | -83.9807
6.5104 | -83.9795
6.5108 | -83.979
6.5111 | -83.9783
6.5114 | -83.9775
6.5117 | -83.9772
6.512 | -83.9768
6.5124 | -83.976
6.5127 | -83.9756
6.513 | -83.9754
6.5133 | -83.9744
6.5137 | -83.9736
6.514 | -83.9726
6.5143 | -83.9715
6.5146 | -83.971
6.5149 | -83.9704
6.5153 | -83.9703
6.5156 | -83.9699
6.5159 | -83.9689
6.5162 | -83.9682
6.5165 | -83.967
6.5169 | -83.9661
6.5172 | -83.9653
6.5175 | -83.9652
6.5178 | -83.9645
6.5181 | -83.9635
6.5185 | -83.9623
6.5188 | -83.9613
6.5191 | -83.9605
6.5194 | -83.9602
6.5198 | -83.9594
6.5201 | -83.959
6.5204 | -83.9583
6.5207 | -83.9576
6.521 | -83.9568
6.5214 | -83.9563
6.5217 | -83.9555
6.522 | -83.9547
6.5223 | -83.9542
6.5226 | -83.9539
6.523 | -83.9533
6.5233 | -83.9521
6.5236 | -83.9506
6.5239 | -83.9498
6.5243 | -83.9504
6.5246 | -83.9508
6.5249 | -83.9502
6.5252 | -83.9503
6.5255 | -83.95
6.5259 | -83.949
6.5262 | -83.9476
6.5265 | -83.9471
6.5268 | -83.947
6.5271 | -83.9465
6.5275 | -83.9456
6.5278 | -83.946
6.5281 | -83.9457
6.5284 | -83.9454
6.5288 | -83.9448
6.5291 | -83.9442
6.5294 | -83.9437
6.5297 | -83.9432
6.53 | -83.9427
6.5304 | -83.9421
6.5307 | -83.9412
6.531 | -83.9404
6.5313 | -83.9401
6.5316 | -83.9399
6.532 | -83.9396
6.5323 | -83.9399
6.5326 | -83.9421
6.5329 | -83.9442
6.5333 | -83.9448
6.5336 | -83.945
6.5339 | -83.9445
6.5342 | -83.9437
6.5345 | -83.9431
6.5349 | -83.9429
6.5352 | -83.9426
6.5355 | -83.9416
6.5358 | -83.9388
6.5361 | -83.9385
6.5365 | -83.9377
6.5368 | -83.937
6.5371 | -83.9364
6.5374 | -83.9358
6.5377 | -83.9347
6.5381 | -83.9335
6.5384 | -83.9325
6.5387 | -83.9319
6.539 | -83.9312
6.5394 | -83.9304
6.5397 | -83.9299
6.54 | -83.9293
6.5403 | -83.9289
6.5406 | -83.9283
6.541 | -83.9282
6.5413 | -83.9287
6.5416 | -83.9282
6.5419 | -83.9279
6.5422 | -83.9275
6.5426 | -83.9268
6.5429 | -83.9264
6.5432 | -83.9259
6.5435 | -83.9253
6.5439 | -83.9249
6.5442 | -83.9244
6.5445 | -83.9232
6.5448 | -83.9228
6.5451 | -83.9229
6.5455 | -83.9234
6.5458 | -83.9249
6.5461 | -83.9256
6.5464 | -83.9248
6.5467 | -83.924
6.5471 | -83.924
6.5474 | -83.9232
6.5477 | -83.9228
6.548 | -83.9223
6.5484 | -83.9217
6.5487 | -83.9217
6.549 | -83.9229
6.5493 | -83.9216
6.5496 | -83.9203
6.55 | -83.9207
6.5503 | -83.9207
6.5506 | -83.9204
6.5509 | -83.92
6.5512 | -83.9193
6.5516 | -83.9191
6.5519 | -83.9187
6.5522 | -83.9183
6.5525 | -83.918
6.5529 | -83.918
6.5532 | -83.9179
6.5535 | -83.9172
6.5538 | -83.9166
6.5541 | -83.9164
6.5545 | -83.9172
6.5548 | -83.9171
6.5551 | -83.9173
6.5554 | -83.9166
6.5557 | -83.9159
6.5561 | -83.9152
6.5564 | -83.915
6.5567 | -83.9145
6.557 | -83.9144
6.5574 | -83.9139
6.5577 | -83.9129
6.558 | -83.912
6.5583 | -83.9111
6.5586 | -83.9108
6.559 | -83.9107
6.5593 | -83.9101
6.5596 | -83.91
6.5599 | -83.9102
6.5602 | -83.9097
6.5606 | -83.9086
6.5609 | -83.9081
6.5612 | -83.9079
6.5615 | -83.9075
6.5618 | -83.9077
6.5622 | -83.9072
6.5625 | -83.9066
6.5628 | -83.9061
6.5631 | -83.9053
6.5635 | -83.9051
6.5638 | -83.9058
6.5641 | -83.9056
6.5644 | -83.9054
6.5647 | -83.9053
6.5651 | -83.9057
6.5654 | -83.9064
6.5657 | -83.9069
6.566 | -83.9066
6.5663 | -83.9064
6.5667 | -83.9064
6.567 | -83.9058
6.5673 | -83.9055
6.5676 | -83.9048
6.568 | -83.904
6.5683 | -83.9038
6.5686 | -83.903
6.5689 | -83.9027
6.5692 | -83.9019
6.5696 | -83.9011
6.5699 | -83.9019
6.5702 | -83.9023
6.5705 | -83.9022
6.5708 | -83.9021
6.5712 | -83.9017
6.5715 | -83.902
6.5718 | -83.9019
6.5721 | -83.902
6.5725 | -83.9025
6.5728 | -83.9024
6.5731 | -83.9031
6.5734 | -83.9031
6.5737 | -83.9036
6.5741 | -83.9035
6.5744 | -83.9034
6.5747 | -83.9031
6.575 | -83.9031
6.5753 | -83.9026
6.5757 | -83.9026
6.576 | -83.9025
6.5763 | -83.9025
6.5766 | -83.9027
6.577 | -83.903
6.5773 | -83.9034
6.5776 | -83.9022
6.5779 | -83.9014
6.5782 | -83.9012
6.5786 | -83.9013
6.5789 | -83.9014
6.5792 | -83.9015
6.5795 | -83.9017
6.5798 | -83.9019
6.5802 | -83.9021
6.5805 | -83.9022
6.5808 | -83.9023
6.5811 | -83.9025
6.5815 | -83.903
6.5818 | -83.9038
6.5821 | -83.9049
6.5824 | -83.9062
6.5827 | -83.9064
6.5831 | -83.9075
6.5834 | -83.9086
6.5837 | -83.9087
6.584 | -83.9088
6.5843 | -83.9088
6.5847 | -83.9086
6.585 | -83.9087
6.5853 | -83.9094
6.5856 | -83.9103
6.5859 | -83.9104
6.5863 | -83.9104
6.5866 | -83.9112
6.5869 | -83.9117
6.5872 | -83.9117
6.5876 | -83.9117
6.5879 | -83.9118
6.5882 | -83.9133
6.5885 | -83.9149
6.5888 | -83.9153
6.5892 | -83.9154
6.5895 | -83.9155
6.5898 | -83.9157
6.5901 | -83.9164
6.5904 | -83.9165
6.5908 | -83.9168
6.5911 | -83.9172
6.5914 | -83.9174
6.5917 | -83.9186
6.5921 | -83.9192
6.5924 | -83.9206
6.5927 | -83.9225
6.593 | -83.9233
6.5933 | -83.9234
6.5937 | -83.9236
6.594 | -83.9236
6.5943 | -83.924
6.5946 | -83.9243
6.5949 | -83.9245
6.5953 | -83.9252
6.5956 | -83.9256
6.5959 | -83.9258
6.5962 | -83.9263
6.5966 | -83.927
6.5969 | -83.9275
6.5972 | -83.9282
6.5975 | -83.9287
6.5978 | -83.9296
6.5982 | -83.9302
6.5985 | -83.9306
6.5988 | -83.9312
6.5991 | -83.9314
6.5994 | -83.9314
6.5998 | -83.9308
6.6001 | -83.9313
6.6004 | -83.9322
6.6007 | -83.9329
6.6011 | -83.9334
6.6014 | -83.9334
6.6017 | -83.934
6.602 | -83.9362
6.6023 | -83.9379
6.6027 | -83.9387
6.603 | -83.9392
6.6033 | -83.9399
6.6036 | -83.9413
6.6039 | -83.9424
6.6043 | -83.943
6.6046 | -83.9437
6.6049 | -83.9445
6.6052 | -83.9449
6.6056 | -83.9453
6.6059 | -83.9458
6.6062 | -83.9465
6.6065 | -83.947
6.6068 | -83.9476
6.6072 | -83.9478
6.6075 | -83.9483
6.6078 | -83.9487
6.6081 | -83.9494
6.6084 | -83.95
6.6088 | -83.95
6.6091 | -83.9504
6.6094 | -83.95
6.6097 | -83.9498
6.61 | -83.9508
6.6104 | -83.9515
6.6107 | -83.9519
6.611 | -83.9532
6.6113 | -83.9547
6.6117 | -83.9558
6.612 | -83.9567
6.6123 | -83.9575
6.6126 | -83.9575
6.6129 | -83.9576
6.6133 | -83.9583
6.6136 | -83.9595
6.6139 | -83.9606
6.6142 | -83.9612
6.6145 | -83.9618
6.6149 | -83.9627
6.6152 | -83.9635
6.6155 | -83.9641
6.6158 | -83.9644
6.6162 | -83.9643
6.6165 | -83.9646
6.6168 | -83.964
6.6171 | -83.9618
6.6174 | -83.9614
6.6178 | -83.9623
6.6181 | -83.9629
6.6184 | -83.9636
6.6187 | -83.9633
6.619 | -83.9619
6.6194 | -83.963
6.6197 | -83.9635
6.62 | -83.9641
6.6203 | -83.9646
6.6207 | -83.9654
6.621 | -83.9662
6.6213 | -83.9674
6.6216 | -83.9682
6.6219 | -83.9689
6.6223 | -83.9696
6.6226 | -83.9704
6.6229 | -83.9713
6.6232 | -83.9716
6.6235 | -83.9721
6.6239 | -83.9728
6.6242 | -83.9735
6.6245 | -83.9743
6.6248 | -83.9752
6.6252 | -83.9769
6.6255 | -83.9781
6.6258 | -83.9788
6.6261 | -83.9795
6.6264 | -83.9805
6.6268 | -83.9813
6.6271 | -83.9821
6.6274 | -83.9821
6.6277 | -83.9822
6.628 | -83.9832
6.6284 | -83.9843
6.6287 | -83.9859
6.629 | -83.9877
6.6293 | -83.9905
6.6296 | -83.9915
6.63 | -83.9923
6.6303 | -83.9929
6.6306 | -83.9945
6.6309 | -83.9958
6.6313 | -83.9968
6.6316 | -83.9975
6.6319 | -83.9982
6.6322 | -83.9991
6.6325 | -83.9999
6.6329 | -84
6.6332 | -84.0001
6.6335 | -84.001
6.6338 | -84.0029
6.6341 | -84.0035
6.6345 | -84.0048
6.6348 | -84.0064
6.6351 | -84.0073
6.6354 | -84.0081
6.6358 | -84.0086
6.6361 | -84.0095
6.6364 | -84.0104
6.6367 | -84.0109
6.637 | -84.0111
6.6374 | -84.0112
6.6377 | -84.0118
6.638 | -84.0128
6.6383 | -84.0142
6.6386 | -84.0161
6.639 | -84.0173
6.6393 | -84.0181
6.6396 | -84.0186
6.6399 | -84.0189
6.6403 | -84.0199
6.6406 | -84.0207
6.6409 | -84.022
6.6412 | -84.0234
6.6415 | -84.023
6.6419 | -84.0221
6.6422 | -84.0224
6.6425 | -84.0232
6.6428 | -84.0239
6.6431 | -84.0248
6.6435 | -84.0253
6.6438 | -84.0259
6.6441 | -84.0267
6.6444 | -84.0278
6.6448 | -84.0293
6.6451 | -84.0298
6.6454 | -84.0302
6.6457 | -84.031
6.646 | -84.032
6.6464 | -84.033
6.6467 | -84.0338
6.647 | -84.0347
6.6473 | -84.036
6.6476 | -84.037
6.648 | -84.0377
6.6483 | -84.0383
6.6486 | -84.0393
6.6489 | -84.0401
6.6493 | -84.0411
6.6496 | -84.0417
6.6499 | -84.0425
6.6502 | -84.0433
6.6505 | -84.0441
6.6509 | -84.0453
6.6512 | -84.0462
6.6515 | -84.0485
6.6518 | -84.0508
6.6521 | -84.0521
6.6525 | -84.0529
6.6528 | -84.054
6.6531 | -84.055
6.6534 | -84.0561
6.6537 | -84.0573
6.6541 | -84.0588
6.6544 | -84.0604
6.6547 | -84.0619
6.655 | -84.0633
6.6554 | -84.0644
6.6557 | -84.0651
6.656 | -84.0658
6.6563 | -84.0668
6.6566 | -84.0686
6.657 | -84.0689
6.6573 | -84.0693
6.6576 | -84.071
6.6579 | -84.0714
6.6582 | -84.0715
6.6586 | -84.0723
6.6589 | -84.0736
6.6592 | -84.0749
6.6595 | -84.0759
6.6599 | -84.0769
6.6602 | -84.078
6.6605 | -84.079
6.6608 | -84.0798
6.6611 | -84.0811
6.6615 | -84.0827
6.6618 | -84.0836
6.6621 | -84.084
6.6624 | -84.0849
6.6627 | -84.086
6.6631 | -84.087
6.6634 | -84.088
6.6637 | -84.0892
6.664 | -84.0904
6.6644 | -84.0914
6.6647 | -84.0924
6.665 | -84.0937
6.6653 | -84.095
6.6656 | -84.0967
6.666 | -84.0976
6.6663 | -84.0986
6.6666 | -84.0996
6.6669 | -84.101
6.6672 | -84.1021
6.6676 | -84.1031
6.6679 | -84.1042
6.6682 | -84.1048
6.6685 | -84.1064
6.6689 | -84.1083
6.6692 | -84.1097
6.6695 | -84.1109
6.6698 | -84.1119
6.6701 | -84.1126
6.6705 | -84.1138
6.6708 | -84.1151
6.6711 | -84.1162
6.6714 | -84.1179
6.6717 | -84.12
6.6721 | -84.1214
6.6724 | -84.1224
6.6727 | -84.1236
6.673 | -84.1246
6.6734 | -84.1253
6.6737 | -84.1265
6.674 | -84.1279
6.6743 | -84.129
6.6746 | -84.1298
6.675 | -84.1311
6.6753 | -84.1312
6.6756 | -84.1314
6.6759 | -84.1327
6.6762 | -84.1338
6.6766 | -84.1348
6.6769 | -84.1356
6.6772 | -84.1368
6.6775 | -84.1376
6.6778 | -84.1381
6.6782 | -84.1393
6.6785 | -84.1402
6.6788 | -84.1411
6.6791 | -84.1427
6.6795 | -84.1439
6.6798 | -84.1447
6.6801 | -84.1455
6.6804 | -84.1463
6.6807 | -84.1475
6.6811 | -84.1487
6.6814 | -84.1501
6.6817 | -84.1514
6.682 | -84.1525
6.6823 | -84.1537
6.6827 | -84.1549
6.683 | -84.156
6.6833 | -84.1572
6.6836 | -84.1584
6.684 | -84.16
6.6843 | -84.1612
6.6846 | -84.1625
6.6849 | -84.1636
6.6852 | -84.1648
6.6856 | -84.166
6.6859 | -84.1667
6.6862 | -84.1681
6.6865 | -84.1697
6.6868 | -84.1705
6.6872 | -84.1716
6.6875 | -84.1728
6.6878 | -84.1739
6.6881 | -84.1755
6.6885 | -84.1771
6.6888 | -84.1782
6.6891 | -84.1782
6.6894 | -84.1784
6.6897 | -84.1796
6.6901 | -84.181
6.6904 | -84.1822
6.6907 | -84.1834
6.691 | -84.1846
6.6913 | -84.1859
6.6917 | -84.1872
6.692 | -84.1882
6.6923 | -84.1889
6.6926 | -84.19
6.693 | -84.1911
6.6933 | -84.1928
6.6936 | -84.194
6.6939 | -84.1956
6.6942 | -84.1969
6.6946 | -84.1982
6.6949 | -84.1995
6.6952 | -84.2011
6.6955 | -84.2022
6.6958 | -84.2038
6.6962 | -84.2054
6.6965 | -84.2066
6.6968 | -84.2071
6.6971 | -84.207
6.6974 | -84.2077
6.6978 | -84.209
6.6981 | -84.2102
6.6984 | -84.2117
6.6987 | -84.2129
6.6991 | -84.214
6.6994 | -84.2155
6.6997 | -84.2181
6.7 | -84.2191
6.7003 | -84.2202
6.7007 | -84.2211
6.701 | -84.2219
6.7013 | -84.2238
6.7016 | -84.2271
6.7019 | -84.2297
6.7023 | -84.2314
6.7026 | -84.2321
6.7029 | -84.2335
6.7032 | -84.2345
6.7036 | -84.2363
6.7039 | -84.2378
6.7042 | -84.2397
6.7045 | -84.2411
6.7048 | -84.2427
6.7052 | -84.2438
6.7055 | -84.2447
6.7058 | -84.2468
6.7061 | -84.2483
6.7064 | -84.249
6.7068 | -84.2496
6.7071 | -84.2507
6.7074 | -84.2517
6.7077 | -84.2524
6.7081 | -84.2527
6.7084 | -84.2536
6.7087 | -84.2551
6.709 | -84.2556
6.7093 | -84.2556
6.7097 | -84.2561
6.71 | -84.2573
6.7103 | -84.2586
6.7106 | -84.2597
6.7109 | -84.26
6.7113 | -84.2605
6.7116 | -84.2618
6.7119 | -84.263
6.7122 | -84.2636
6.7126 | -84.264
6.7129 | -84.2647
6.7132 | -84.2659
6.7135 | -84.2672
6.7138 | -84.2681
6.7142 | -84.2688
6.7145 | -84.2694
6.7148 | -84.271
6.7151 | -84.2715
6.7154 | -84.2725
6.7158 | -84.2734
6.7161 | -84.2753
6.7164 | -84.2776
6.7167 | -84.2785
6.7171 | -84.2788
6.7174 | -84.2796
6.7177 | -84.2805
6.718 | -84.2816
6.7183 | -84.2827
6.7187 | -84.2838
6.719 | -84.2848
6.7193 | -84.2857
6.7196 | -84.2862
6.7199 | -84.2873
6.7203 | -84.2885
6.7206 | -84.2895
6.7209 | -84.2902
6.7212 | -84.291
6.7215 | -84.2923
6.7219 | -84.2934
6.7222 | -84.294
6.7225 | -84.2948
6.7228 | -84.2959
6.7232 | -84.297
6.7235 | -84.2978
6.7238 | -84.2983
6.7241 | -84.2992
6.7244 | -84.3003
6.7248 | -84.3022
6.7251 | -84.3042
6.7254 | -84.3053
6.7257 | -84.3058
6.726 | -84.3065
6.7264 | -84.3072
6.7267 | -84.3085
6.727 | -84.3088
6.7273 | -84.3086
6.7277 | -84.3094
6.728 | -84.3106
6.7283 | -84.3116
6.7286 | -84.3131
6.7289 | -84.3141
6.7293 | -84.3149
6.7296 | -84.3159
6.7299 | -84.317
6.7302 | -84.3177
6.7305 | -84.3183
6.7309 | -84.3186
6.7312 | -84.3193
6.7315 | -84.3202
6.7318 | -84.3214
6.7322 | -84.3233
6.7325 | -84.3248
6.7328 | -84.3256
6.7331 | -84.3259
6.7334 | -84.3271
6.7338 | -84.3279
6.7341 | -84.3288
6.7344 | -84.3303
6.7347 | -84.3312
6.735 | -84.3321
6.7354 | -84.3328
6.7357 | -84.3338
6.736 | -84.3345
6.7363 | -84.3354
6.7367 | -84.3361
6.737 | -84.3367
6.7373 | -84.3376
6.7376 | -84.3384
6.7379 | -84.3391
6.7383 | -84.3397
6.7386 | -84.3404
6.7389 | -84.3413
6.7392 | -84.3422
6.7395 | -84.3432
6.7399 | -84.3437
6.7402 | -84.344
6.7405 | -84.3437
6.7408 | -84.3442
6.7412 | -84.3451
6.7415 | -84.3455
6.7418 | -84.3453
6.7421 | -84.3456
6.7424 | -84.3472
6.7428 | -84.3491
6.7431 | -84.3502
6.7434 | -84.3513
6.7437 | -84.3521
6.744 | -84.3532
6.7444 | -84.354
6.7447 | -84.3547
6.745 | -84.3554
6.7453 | -84.3559
6.7456 | -84.3573
6.746 | -84.3589
6.7463 | -84.3612
6.7466 | -84.3642
6.7469 | -84.3654
6.7473 | -84.3665
6.7476 | -84.3678
6.7479 | -84.3685
6.7482 | -84.3691
6.7485 | -84.3702
6.7489 | -84.3718
6.7492 | -84.3725
6.7495 | -84.3734
6.7498 | -84.3739
6.7501 | -84.3749
6.7505 | -84.3763
6.7508 | -84.377
6.7511 | -84.378
6.7514 | -84.3796
6.7518 | -84.3814
6.7521 | -84.3835
6.7524 | -84.3849
6.7527 | -84.3856
6.753 | -84.3861
6.7534 | -84.3872
6.7537 | -84.3879
6.754 | -84.3877
6.7543 | -84.388
6.7546 | -84.3887
6.755 | -84.3894
6.7553 | -84.3898
6.7556 | -84.3914
6.7559 | -84.3932
6.7563 | -84.3937
6.7566 | -84.3946
6.7569 | -84.395
6.7572 | -84.3952
6.7575 | -84.396
6.7579 | -84.3975
6.7582 | -84.3986
6.7585 | -84.3987
6.7588 | -84.3992
6.7591 | -84.3996
6.7595 | -84.4005
6.7598 | -84.4015
6.7601 | -84.4023
6.7604 | -84.4034
6.7608 | -84.404
6.7611 | -84.4032
6.7614 | -84.4032
6.7617 | -84.4038
6.762 | -84.4044
6.7624 | -84.4051
6.7627 | -84.4055
6.763 | -84.4062
6.7633 | -84.407
6.7636 | -84.4074
6.764 | -84.4083
6.7643 | -84.4091
6.7646 | -84.4098
6.7649 | -84.411
6.7652 | -84.4127
6.7656 | -84.4141
6.7659 | -84.4148
6.7662 | -84.4156
6.7665 | -84.4164
6.7669 | -84.4172
6.7672 | -84.4179
6.7675 | -84.4184
6.7678 | -84.4176
6.7681 | -84.4179
6.7685 | -84.4186
6.7688 | -84.4193
6.7691 | -84.42
6.7694 | -84.4209
6.7697 | -84.4214
6.7701 | -84.4217
6.7704 | -84.4218
6.7707 | -84.4223
6.771 | -84.4229
6.7714 | -84.4242
6.7717 | -84.4249
6.772 | -84.4257
6.7723 | -84.4269
6.7726 | -84.4286
6.773 | -84.4292
6.7733 | -84.4311
6.7736 | -84.4325
6.7739 | -84.4332
6.7742 | -84.434
6.7746 | -84.4355
6.7749 | -84.4358
6.7752 | -84.4366
6.7755 | -84.4385
6.7759 | -84.4402
6.7762 | -84.4407
6.7765 | -84.4415
6.7768 | -84.4426
6.7771 | -84.4432
6.7775 | -84.444
6.7778 | -84.4446
6.7781 | -84.4442
6.7784 | -84.4447
6.7787 | -84.4462
6.7791 | -84.4474
6.7794 | -84.4482
6.7797 | -84.4487
6.78 | -84.4493
6.7804 | -84.4497
6.7807 | -84.4495
6.781 | -84.4486
6.7813 | -84.4484
6.7816 | -84.4493
6.782 | -84.4503
6.7823 | -84.4511
6.7826 | -84.4513
6.7829 | -84.4517
6.7832 | -84.4529
6.7836 | -84.4533
6.7839 | -84.4543
6.7842 | -84.4569
6.7845 | -84.4596
6.7849 | -84.4602
6.7852 | -84.46
6.7855 | -84.4604
6.7858 | -84.4609
6.7861 | -84.4605
6.7865 | -84.4614
6.7868 | -84.4628
6.7871 | -84.4641
6.7874 | -84.4647
6.7877 | -84.4659
6.7881 | -84.4666
6.7884 | -84.4661
6.7887 | -84.4674
6.789 | -84.4694
6.7893 | -84.4701
6.7897 | -84.4704
6.79 | -84.4706
6.7903 | -84.471
6.7906 | -84.472
6.791 | -84.4735
6.7913 | -84.4744
6.7916 | -84.4746
6.7919 | -84.475
6.7922 | -84.4752
6.7926 | -84.4753
6.7929 | -84.4761
6.7932 | -84.4764
6.7935 | -84.4766
6.7938 | -84.4771
6.7942 | -84.4778
6.7945 | -84.4783
6.7948 | -84.4786
6.7951 | -84.4789
6.7955 | -84.4781
6.7958 | -84.4785
6.7961 | -84.4792
6.7964 | -84.4794
6.7967 | -84.4799
6.7971 | -84.4811
6.7974 | -84.4825
6.7977 | -84.4828
6.798 | -84.4828
6.7983 | -84.4835
6.7987 | -84.4834
6.799 | -84.4825
6.7993 | -84.4825
6.7996 | -84.4829
6.8 | -84.4835
6.8003 | -84.4841
6.8006 | -84.485
6.8009 | -84.4857
6.8012 | -84.4862
6.8016 | -84.4866
6.8019 | -84.4875
6.8022 | -84.4885
6.8025 | -84.489
6.8028 | -84.4897
6.8032 | -84.4903
6.8035 | -84.4911
6.8038 | -84.4909
6.8041 | -84.4907
6.8045 | -84.4913
6.8048 | -84.4912
6.8051 | -84.4915
6.8054 | -84.4913
6.8057 | -84.4915
6.8061 | -84.4922
6.8064 | -84.4924
6.8067 | -84.4927
6.807 | -84.4933
6.8073 | -84.4936
6.8077 | -84.4938
6.808 | -84.494
6.8083 | -84.4934
6.8086 | -84.4933
6.809 | -84.4931
6.8093 | -84.4934
6.8096 | -84.4936
6.8099 | -84.4937
6.8102 | -84.4938
6.8106 | -84.494
6.8109 | -84.4941
6.8112 | -84.4944
6.8115 | -84.4947
6.8118 | -84.495
6.8122 | -84.4957
6.8125 | -84.4967
6.8128 | -84.4972
6.8131 | -84.4983
6.8134 | -84.4984
6.8138 | -84.4989
6.8141 | -84.4995
6.8144 | -84.5004
6.8147 | -84.5006
6.8151 | -84.5012
6.8154 | -84.5019
6.8157 | -84.5024
6.816 | -84.5023
6.8163 | -84.5021
6.8167 | -84.5024
6.817 | -84.5019
6.8173 | -84.5019
6.8176 | -84.5019
6.8179 | -84.5022
6.8183 | -84.5024
6.8186 | -84.5025
6.8189 | -84.5024
6.8192 | -84.5027
6.8196 | -84.5028
6.8199 | -84.5024
6.8202 | -84.5019
6.8205 | -84.5023
6.8208 | -84.5026
6.8212 | -84.5024
6.8215 | -84.5028
6.8218 | -84.5034
6.8221 | -84.5036
6.8224 | -84.5038
6.8228 | -84.5042
6.8231 | -84.5047
6.8234 | -84.5052
6.8237 | -84.5056
6.8241 | -84.5061
6.8244 | -84.5065
6.8247 | -84.5071
6.825 | -84.5076
6.8253 | -84.5078
6.8257 | -84.508
6.826 | -84.5081
6.8263 | -84.5083
6.8266 | -84.5087
6.8269 | -84.509
6.8273 | -84.5089
6.8276 | -84.509
6.8279 | -84.5092
6.8282 | -84.5104
6.8286 | -84.5114
6.8289 | -84.5117
6.8292 | -84.5124
6.8295 | -84.5131
6.8298 | -84.513
6.8302 | -84.5134
6.8305 | -84.5142
6.8308 | -84.5151
6.8311 | -84.5155
6.8314 | -84.5159
6.8318 | -84.5167
6.8321 | -84.5179
6.8324 | -84.5185
6.8327 | -84.5194
6.8331 | -84.5198
6.8334 | -84.5198
6.8337 | -84.5203
6.834 | -84.5206
6.8343 | -84.521
6.8347 | -84.5227
6.835 | -84.5239
6.8353 | -84.5246
6.8356 | -84.5254
6.8359 | -84.5261
6.8363 | -84.5269
6.8366 | -84.5282
6.8369 | -84.5294
6.8372 | -84.5304
6.8375 | -84.5318
6.8379 | -84.5325
6.8382 | -84.5328
6.8385 | -84.5332
6.8388 | -84.5335
6.8392 | -84.5341
6.8395 | -84.5347
6.8398 | -84.5349
6.8401 | -84.5358
6.8404 | -84.5366
6.8408 | -84.538
6.8411 | -84.5389
6.8414 | -84.5394
6.8417 | -84.54
6.842 | -84.5405
6.8424 | -84.5409
6.8427 | -84.5413
6.843 | -84.5413
6.8433 | -84.5418
6.8437 | -84.5419
6.844 | -84.5432
6.8443 | -84.5446
6.8446 | -84.546
6.8449 | -84.5467
6.8453 | -84.5472
6.8456 | -84.5488
6.8459 | -84.5504
6.8462 | -84.5503
6.8465 | -84.5503
6.8469 | -84.5506
6.8472 | -84.5504
6.8475 | -84.5503
6.8478 | -84.5516
6.8482 | -84.5531
6.8485 | -84.5534
6.8488 | -84.5536
6.8491 | -84.5539
6.8494 | -84.555
6.8498 | -84.5554
6.8501 | -84.5561
6.8504 | -84.5575
6.8507 | -84.5595
6.851 | -84.5605
6.8514 | -84.5622
6.8517 | -84.5633
6.852 | -84.5636
6.8523 | -84.5636
6.8527 | -84.5635
6.853 | -84.5639
6.8533 | -84.5641
6.8536 | -84.5644
6.8539 | -84.5648
6.8543 | -84.5657
6.8546 | -84.5678
6.8549 | -84.5694
6.8552 | -84.5692
6.8555 | -84.5708
6.8559 | -84.5734
6.8562 | -84.5749
6.8565 | -84.5754
6.8568 | -84.5757
6.8571 | -84.5761
6.8575 | -84.5769
6.8578 | -84.5775
6.8581 | -84.5779
6.8584 | -84.5785
6.8588 | -84.5792
6.8591 | -84.5791
6.8594 | -84.5789
6.8597 | -84.579
6.86 | -84.5791
6.8604 | -84.5799
6.8607 | -84.5799
6.861 | -84.5801
6.8613 | -84.5805
6.8616 | -84.5809
6.862 | -84.5812
6.8623 | -84.5819
6.8626 | -84.5828
6.8629 | -84.5837
6.8633 | -84.5849
6.8636 | -84.5856
6.8639 | -84.5865
6.8642 | -84.5876
6.8645 | -84.588
6.8649 | -84.5882
6.8652 | -84.5881
6.8655 | -84.5882
6.8658 | -84.5878
6.8661 | -84.5873
6.8665 | -84.5875
6.8668 | -84.5879
6.8671 | -84.5882
6.8674 | -84.5891
6.8678 | -84.5893
6.8681 | -84.5894
6.8684 | -84.5897
6.8687 | -84.5899
6.869 | -84.5903
6.8694 | -84.5906
6.8697 | -84.5912
6.87 | -84.5919
6.8703 | -84.593
6.8706 | -84.5947
6.871 | -84.595
6.8713 | -84.5959
6.8716 | -84.5971
6.8719 | -84.5977
6.8723 | -84.599
6.8726 | -84.5998
6.8729 | -84.5994
6.8732 | -84.5994
6.8735 | -84.5998
6.8739 | -84.6002
6.8742 | -84.6003
6.8745 | -84.6006
6.8748 | -84.6007
6.8751 | -84.6009
6.8755 | -84.6012
6.8758 | -84.602
6.8761 | -84.6021
6.8764 | -84.6025
6.8768 | -84.6033
6.8771 | -84.6043
6.8774 | -84.6043
6.8777 | -84.6044
6.878 | -84.6046
6.8784 | -84.6048
6.8787 | -84.6044
6.879 | -84.6053
6.8793 | -84.6054
6.8796 | -84.6058
6.88 | -84.6054
6.8803 | -84.6046
6.8806 | -84.6049
6.8809 | -84.6054
6.8812 | -84.6057
6.8816 | -84.6058
6.8819 | -84.6061
6.8822 | -84.6064
6.8825 | -84.6065
6.8829 | -84.6069
6.8832 | -84.6079
6.8835 | -84.6086
6.8838 | -84.609
6.8841 | -84.6095
6.8845 | -84.6097
6.8848 | -84.6102
6.8851 | -84.6102
6.8854 | -84.6088
6.8857 | -84.6081
6.8861 | -84.6078
6.8864 | -84.6079
6.8867 | -84.6089
6.887 | -84.6099
6.8874 | -84.6105
6.8877 | -84.611
6.888 | -84.6115
6.8883 | -84.6125
6.8886 | -84.6129
6.889 | -84.6132
6.8893 | -84.6134
6.8896 | -84.6134
6.8899 | -84.6132
6.8902 | -84.6144
6.8906 | -84.6153
6.8909 | -84.6158
6.8912 | -84.6162
6.8915 | -84.6168
6.8919 | -84.6168
6.8922 | -84.6164
6.8925 | -84.6169
6.8928 | -84.6175
6.8931 | -84.6178
6.8935 | -84.6178
6.8938 | -84.619
6.8941 | -84.621
6.8944 | -84.6217
6.8947 | -84.6217
6.8951 | -84.6224
6.8954 | -84.6227
6.8957 | -84.6229
6.896 | -84.6233
6.8964 | -84.6236
6.8967 | -84.6235
6.897 | -84.6238
6.8973 | -84.6239
6.8976 | -84.6247
6.898 | -84.6251
6.8983 | -84.6257
6.8986 | -84.6267
6.8989 | -84.6271
6.8992 | -84.6275
6.8996 | -84.6278
6.8999 | -84.6283
6.9002 | -84.6293
6.9005 | -84.6303
6.9009 | -84.6319
6.9012 | -84.6339
6.9015 | -84.6341
6.9018 | -84.6347
6.9021 | -84.6351
6.9025 | -84.6355
6.9028 | -84.6359
6.9031 | -84.6372
6.9034 | -84.638
6.9037 | -84.6385
6.9041 | -84.639
6.9044 | -84.6392
6.9047 | -84.6395
6.905 | -84.6403
6.9053 | -84.6405
6.9057 | -84.6413
6.906 | -84.6422
6.9063 | -84.6426
6.9066 | -84.6433
6.907 | -84.6431
6.9073 | -84.6438
6.9076 | -84.6442
6.9079 | -84.6447
6.9082 | -84.6448
6.9086 | -84.6449
6.9089 | -84.6452
6.9092 | -84.6449
6.9095 | -84.6445
6.9098 | -84.6446
6.9102 | -84.6455
6.9105 | -84.6461
6.9108 | -84.6462
6.9111 | -84.6464
6.9115 | -84.6467
6.9118 | -84.647
6.9121 | -84.6474
6.9124 | -84.6473
6.9127 | -84.6459
6.9131 | -84.646
6.9134 | -84.6461
6.9137 | -84.6469
6.914 | -84.6482
6.9143 | -84.6487
6.9147 | -84.6493
6.915 | -84.6496
6.9153 | -84.6501
6.9156 | -84.6504
6.916 | -84.6514
6.9163 | -84.652
6.9166 | -84.6528
6.9169 | -84.6537
6.9172 | -84.6547
6.9176 | -84.6548
6.9179 | -84.6558
6.9182 | -84.6569
6.9185 | -84.6573
6.9188 | -84.6574
6.9192 | -84.6575
6.9195 | -84.6579
6.9198 | -84.6586
6.9201 | -84.6591
6.9205 | -84.6598
6.9208 | -84.6607
6.9211 | -84.6613
6.9214 | -84.6618
6.9217 | -84.6623
6.9221 | -84.6628
6.9224 | -84.6631
6.9227 | -84.6635
6.923 | -84.6641
6.9233 | -84.6646
6.9237 | -84.6653
6.924 | -84.6658
6.9243 | -84.6662
6.9246 | -84.6674
6.925 | -84.669
6.9253 | -84.6702
6.9256 | -84.6715
6.9259 | -84.6719
6.9262 | -84.6722
6.9266 | -84.6726
6.9269 | -84.6734
6.9272 | -84.6746
6.9275 | -84.6751
6.9278 | -84.6762
6.9282 | -84.6771
6.9285 | -84.6781
6.9288 | -84.6789
6.9291 | -84.6793
6.9294 | -84.6798
6.9298 | -84.6806
6.9301 | -84.6812
6.9304 | -84.6818
6.9307 | -84.6827
6.9311 | -84.6831
6.9314 | -84.6835
6.9317 | -84.6841
6.932 | -84.685
6.9323 | -84.6861
6.9327 | -84.687
6.933 | -84.6879
6.9333 | -84.6885
6.9336 | -84.6893
6.9339 | -84.6897
6.9343 | -84.6913
6.9346 | -84.6919
6.9349 | -84.693
6.9352 | -84.6938
6.9356 | -84.6945
6.9359 | -84.6952
6.9362 | -84.6959
6.9365 | -84.6965
6.9368 | -84.6969
6.9372 | -84.6976
6.9375 | -84.6985
6.9378 | -84.6993
6.9381 | -84.6996
6.9384 | -84.6995
6.9388 | -84.7005
6.9391 | -84.702
6.9394 | -84.7032
6.9397 | -84.7036
6.9401 | -84.7054
6.9404 | -84.707
6.9407 | -84.7085
6.941 | -84.7094
6.9413 | -84.7101
6.9417 | -84.7108
6.942 | -84.7116
6.9423 | -84.7125
6.9426 | -84.713
6.9429 | -84.7137
6.9433 | -84.715
6.9436 | -84.7163
6.9439 | -84.7177
6.9442 | -84.7186
6.9446 | -84.7186
6.9449 | -84.7192
6.9452 | -84.7204
6.9455 | -84.7221
6.9458 | -84.7231
6.9462 | -84.7237
6.9465 | -84.7243
6.9468 | -84.7254
6.9471 | -84.7264
6.9474 | -84.727
6.9478 | -84.7281
6.9481 | -84.7291
6.9484 | -84.7298
6.9487 | -84.7306
6.949 | -84.7314
6.9494 | -84.7325
6.9497 | -84.7337
6.95 | -84.7344
6.9503 | -84.7355
6.9507 | -84.7372
6.951 | -84.7387
6.9513 | -84.7392
6.9516 | -84.7398
6.9519 | -84.7406
6.9523 | -84.7397
6.9526 | -84.7387
6.9529 | -84.7401
6.9532 | -84.7408
6.9535 | -84.7418
6.9539 | -84.7427
6.9542 | -84.7435
6.9545 | -84.7444
6.9548 | -84.7455
6.9552 | -84.7466
6.9555 | -84.7472
6.9558 | -84.7481
6.9561 | -84.7487
6.9564 | -84.7501
6.9568 | -84.7519
6.9571 | -84.7525
6.9574 | -84.7535
6.9577 | -84.7551
6.958 | -84.7571
6.9584 | -84.7582
6.9587 | -84.7587
6.959 | -84.7599
6.9593 | -84.761
6.9597 | -84.7616
6.96 | -84.7627
6.9603 | -84.763
6.9606 | -84.7631
6.9609 | -84.7641
6.9613 | -84.7652
6.9616 | -84.766
6.9619 | -84.7671
6.9622 | -84.7682
6.9625 | -84.7698
6.9629 | -84.7704
6.9632 | -84.7706
6.9635 | -84.7719
6.9638 | -84.7726
6.9642 | -84.773
6.9645 | -84.7736
6.9648 | -84.7748
6.9651 | -84.7758
6.9654 | -84.7766
6.9658 | -84.7782
6.9661 | -84.7793
6.9664 | -84.7801
6.9667 | -84.7814
6.967 | -84.783
6.9674 | -84.7843
6.9677 | -84.785
6.968 | -84.7859
6.9683 | -84.787
6.9687 | -84.7881
6.969 | -84.7893
6.9693 | -84.7902
6.9696 | -84.7912
6.9699 | -84.792
6.9703 | -84.7929
6.9706 | -84.7938
6.9709 | -84.7949
6.9712 | -84.7959
6.9715 | -84.7969
6.9719 | -84.7982
6.9722 | -84.7993
6.9725 | -84.8
6.9728 | -84.8009
6.9732 | -84.8025
6.9735 | -84.8046
6.9738 | -84.8057
6.9741 | -84.8065
6.9744 | -84.8072
6.9748 | -84.8076
6.9751 | -84.8087
6.9754 | -84.8093
6.9757 | -84.8102
6.976 | -84.8114
6.9764 | -84.8124
6.9767 | -84.813
6.977 | -84.8138
6.9773 | -84.8148
6.9776 | -84.8158
6.978 | -84.8162
6.9783 | -84.8158
6.9786 | -84.8159
6.9789 | -84.8166
6.9793 | -84.8177
6.9796 | -84.8183
6.9799 | -84.8188
6.9802 | -84.8198
6.9805 | -84.8206
6.9809 | -84.8217
6.9812 | -84.8226
6.9815 | -84.8235
6.9818 | -84.8241
6.9821 | -84.825
6.9825 | -84.8266
6.9828 | -84.8279
6.9831 | -84.8287
6.9834 | -84.8295
6.9838 | -84.8304
6.9841 | -84.8314
6.9844 | -84.8327
6.9847 | -84.8342
6.985 | -84.8362
6.9854 | -84.8372
6.9857 | -84.8384
6.986 | -84.8395
6.9863 | -84.8401
6.9866 | -84.8408
6.987 | -84.8425
6.9873 | -84.8445
6.9876 | -84.8459
6.9879 | -84.8469
6.9883 | -84.8474
6.9886 | -84.8485
6.9889 | -84.8494
6.9892 | -84.8506
6.9895 | -84.8516
6.9899 | -84.8527
6.9902 | -84.8542
6.9905 | -84.855
6.9908 | -84.8562
6.9911 | -84.8575
6.9915 | -84.8585
6.9918 | -84.8596
6.9921 | -84.8608
6.9924 | -84.8626
6.9928 | -84.8646
6.9931 | -84.8658
6.9934 | -84.8669
6.9937 | -84.8675
6.994 | -84.8683
6.9944 | -84.8693
6.9947 | -84.8706
6.995 | -84.8722
6.9953 | -84.8737
6.9956 | -84.8746
6.996 | -84.8754
6.9963 | -84.8767
6.9966 | -84.8773
6.9969 | -84.8789
6.9972 | -84.8798
6.9976 | -84.8815
6.9979 | -84.8827
6.9982 | -84.8839
6.9985 | -84.885
6.9989 | -84.886
6.9992 | -84.8867
6.9995 | -84.8877
6.9998 | -84.889
7.0001 | -84.8904
7.0005 | -84.8917
7.0008 | -84.8929
7.0011 | -84.8943
7.0014 | -84.8952
7.0017 | -84.8963
7.0021 | -84.8977
7.0024 | -84.8989
7.0027 | -84.8999
7.003 | -84.901
7.0034 | -84.9016
7.0037 | -84.9014
7.004 | -84.9018
7.0043 | -84.903
7.0046 | -84.9039
7.005 | -84.9052
7.0053 | -84.9065
7.0056 | -84.9103
7.0059 | -84.9125
7.0062 | -84.914
7.0066 | -84.9156
7.0069 | -84.9169
7.0072 | -84.9181
7.0075 | -84.9194
7.0079 | -84.9213
7.0082 | -84.9231
7.0085 | -84.9241
7.0088 | -84.9253
7.0091 | -84.9266
7.0095 | -84.9277
7.0098 | -84.9289
7.0101 | -84.93
7.0104 | -84.9311
7.0107 | -84.9322
7.0111 | -84.9335
7.0114 | -84.9337
7.0117 | -84.9345
7.012 | -84.9349
7.0124 | -84.9356
7.0127 | -84.9367
7.013 | -84.9381
7.0133 | -84.9389
7.0136 | -84.9399
7.014 | -84.9409
7.0143 | -84.942
7.0146 | -84.9433
7.0149 | -84.9449
7.0152 | -84.9462
7.0156 | -84.9475
7.0159 | -84.9494
7.0162 | -84.9509
7.0165 | -84.9523
7.0168 | -84.9536
7.0172 | -84.9555
7.0175 | -84.9574
7.0178 | -84.9577
7.0181 | -84.9592
7.0185 | -84.961
7.0188 | -84.9624
7.0191 | -84.963
7.0194 | -84.9636
7.0197 | -84.9645
7.0201 | -84.9665
7.0204 | -84.9697
7.0207 | -84.9726
7.021 | -84.9739
7.0213 | -84.9751
7.0217 | -84.9769
7.022 | -84.9785
7.0223 | -84.98
7.0226 | -84.9811
7.023 | -84.9823
7.0233 | -84.9836
7.0236 | -84.9849
7.0239 | -84.9858
7.0242 | -84.9872
7.0246 | -84.989
7.0249 | -84.9903
7.0252 | -84.9913
7.0255 | -84.9924
7.0258 | -84.9936
7.0262 | -84.9939
7.0265 | -84.9943
7.0268 | -84.9959
7.0271 | -84.9977
7.0275 | -84.9997
7.0278 | -85.0017
7.0281 | -85.003
7.0284 | -85.0046
7.0287 | -85.006
7.0291 | -85.0077
7.0294 | -85.0097
7.0297 | -85.0122
7.03 | -85.0142
7.0303 | -85.0153
7.0307 | -85.0164
7.031 | -85.018
7.0313 | -85.0194
7.0316 | -85.02
7.032 | -85.0217
7.0323 | -85.0246
7.0326 | -85.0265
7.0329 | -85.028
7.0332 | -85.0297
7.0336 | -85.0312
7.0339 | -85.0332
7.0342 | -85.0355
7.0345 | -85.0381
7.0348 | -85.0403
7.0352 | -85.0416
7.0355 | -85.0434
7.0358 | -85.0447
7.0361 | -85.0457
7.0365 | -85.0468
7.0368 | -85.0489
7.0371 | -85.0501
7.0374 | -85.0511
7.0377 | -85.052
7.0381 | -85.053
7.0384 | -85.0544
7.0387 | -85.0556
7.039 | -85.0569
7.0393 | -85.0587
7.0397 | -85.0601
7.04 | -85.0624
7.0403 | -85.0645
7.0406 | -85.0663
7.0409 | -85.0687
7.0413 | -85.0704
7.0416 | -85.0723
7.0419 | -85.0749
7.0422 | -85.0766
7.0426 | -85.078
7.0429 | -85.0793
7.0432 | -85.0808
7.0435 | -85.0823
7.0438 | -85.083
7.0442 | -85.0847
7.0445 | -85.0867
7.0448 | -85.0878
7.0451 | -85.0886
7.0454 | -85.0901
7.0458 | -85.0906
7.0461 | -85.0922
7.0464 | -85.0938
7.0467 | -85.0952
7.0471 | -85.097
7.0474 | -85.099
7.0477 | -85.1004
7.048 | -85.1012
7.0483 | -85.1029
7.0487 | -85.1046
7.049 | -85.1062
7.0493 | -85.1073
7.0496 | -85.1083
7.0499 | -85.1096
7.0503 | -85.1114
7.0506 | -85.1143
7.0509 | -85.1157
7.0512 | -85.1171
7.0516 | -85.1187
7.0519 | -85.1201
7.0522 | -85.1216
7.0525 | -85.1229
7.0528 | -85.1239
7.0532 | -85.1251
7.0535 | -85.1269
7.0538 | -85.1284
7.0541 | -85.1297
7.0544 | -85.1304
7.0548 | -85.1307
7.0551 | -85.1326
7.0554 | -85.1341
7.0557 | -85.1355
7.0561 | -85.1368
7.0564 | -85.1372
7.0567 | -85.1387
7.057 | -85.1409
7.0573 | -85.1423
7.0577 | -85.1434
7.058 | -85.1452
7.0583 | -85.147
7.0586 | -85.1485
7.0589 | -85.1501
7.0593 | -85.1516
7.0596 | -85.1532
7.0599 | -85.1544
7.0602 | -85.1555
7.0606 | -85.1556
7.0609 | -85.1564
7.0612 | -85.1577
7.0615 | -85.159
7.0618 | -85.1604
7.0622 | -85.1615
7.0625 | -85.1628
7.0628 | -85.1643
7.0631 | -85.1662
7.0634 | -85.1675
7.0638 | -85.169
7.0641 | -85.1708
7.0644 | -85.1713
7.0647 | -85.1724
7.0651 | -85.1738
7.0654 | -85.1754
7.0657 | -85.1766
7.066 | -85.1784
7.0663 | -85.1799
7.0667 | -85.1812
7.067 | -85.1828
7.0673 | -85.1842
7.0676 | -85.186
7.0679 | -85.1881
7.0683 | -85.1898
7.0686 | -85.1914
7.0689 | -85.1926
7.0692 | -85.1943
7.0695 | -85.1956
7.0699 | -85.1971
7.0702 | -85.198
7.0705 | -85.1986
7.0708 | -85.2001
7.0712 | -85.2015
7.0715 | -85.203
7.0718 | -85.2052
7.0721 | -85.2072
7.0724 | -85.2087
7.0728 | -85.2098
7.0731 | -85.2108
7.0734 | -85.2122
7.0737 | -85.2133
7.074 | -85.2151
7.0744 | -85.2176
7.0747 | -85.2188
7.075 | -85.2202
7.0753 | -85.2218
7.0757 | -85.2241
7.076 | -85.2256
7.0763 | -85.2276
7.0766 | -85.2293
7.0769 | -85.23
7.0773 | -85.2311
7.0776 | -85.2318
7.0779 | -85.2326
7.0782 | -85.2338
7.0785 | -85.235
7.0789 | -85.2366
7.0792 | -85.2389
7.0795 | -85.241
7.0798 | -85.2425
7.0802 | -85.2442
7.0805 | -85.2456
7.0808 | -85.247
7.0811 | -85.2486
7.0814 | -85.25
7.0818 | -85.2512
7.0821 | -85.2525
7.0824 | -85.2526
7.0827 | -85.2531
7.083 | -85.2546
7.0834 | -85.2559
7.0837 | -85.2581
7.084 | -85.2609
7.0843 | -85.263
7.0847 | -85.2655
7.085 | -85.2678
7.0853 | -85.2696
7.0856 | -85.2707
7.0859 | -85.2716
7.0863 | -85.2725
7.0866 | -85.2737
7.0869 | -85.2746
7.0872 | -85.276
7.0875 | -85.2775
7.0879 | -85.279
7.0882 | -85.2808
7.0885 | -85.2826
7.0888 | -85.2838
7.0891 | -85.2851
7.0895 | -85.2867
7.0898 | -85.2875
7.0901 | -85.289
7.0904 | -85.2901
7.0908 | -85.2912
7.0911 | -85.2928
7.0914 | -85.2939
7.0917 | -85.295
7.092 | -85.297
7.0924 | -85.2984
7.0927 | -85.3014
7.093 | -85.3058
7.0933 | -85.3075
7.0936 | -85.3085
7.094 | -85.31
7.0943 | -85.3111
7.0946 | -85.3139
7.0949 | -85.3177
7.0953 | -85.3192
7.0956 | -85.3203
7.0959 | -85.3216
7.0962 | -85.3229
7.0965 | -85.3241
7.0969 | -85.3253
7.0972 | -85.3264
7.0975 | -85.3277
7.0978 | -85.3288
7.0981 | -85.33
7.0985 | -85.331
7.0988 | -85.3321
7.0991 | -85.3334
7.0994 | -85.3353
7.0998 | -85.3373
7.1001 | -85.3383
7.1004 | -85.3395
7.1007 | -85.341
7.101 | -85.3424
7.1014 | -85.3438
7.1017 | -85.3451
7.102 | -85.3465
7.1023 | -85.3473
7.1026 | -85.3482
7.103 | -85.3495
7.1033 | -85.35
7.1036 | -85.3513
7.1039 | -85.3535
7.1043 | -85.3561
7.1046 | -85.3575
7.1049 | -85.3586
7.1052 | -85.3598
7.1055 | -85.3613
7.1059 | -85.3626
7.1062 | -85.3638
7.1065 | -85.3651
7.1068 | -85.3668
7.1071 | -85.3692
7.1075 | -85.3706
7.1078 | -85.3717
7.1081 | -85.373
7.1084 | -85.3743
7.1087 | -85.3754
7.1091 | -85.3766
7.1094 | -85.3783
7.1097 | -85.3807
7.11 | -85.3827
7.1104 | -85.3842
7.1107 | -85.3854
7.111 | -85.3869
7.1113 | -85.3886
7.1116 | -85.3896
7.112 | -85.3902
7.1123 | -85.3907
7.1126 | -85.3903
7.1129 | -85.3909
7.1132 | -85.3921
7.1136 | -85.3932
7.1139 | -85.3947
7.1142 | -85.397
7.1145 | -85.3995
7.1149 | -85.4013
7.1152 | -85.4033
7.1155 | -85.4048
7.1158 | -85.4062
7.1161 | -85.4062
7.1165 | -85.4068
7.1168 | -85.4079
7.1171 | -85.4097
7.1174 | -85.4118
7.1177 | -85.413
7.1181 | -85.4144
7.1184 | -85.4156
7.1187 | -85.4173
7.119 | -85.42
7.1194 | -85.4234
7.1197 | -85.4253
7.12 | -85.426
7.1203 | -85.4272
7.1206 | -85.4284
7.121 | -85.4293
7.1213 | -85.431
7.1216 | -85.4326
7.1219 | -85.4337
7.1222 | -85.4348
7.1226 | -85.436
7.1229 | -85.4375
7.1232 | -85.439
7.1235 | -85.4402
7.1239 | -85.4411
7.1242 | -85.4422
7.1245 | -85.4435
7.1248 | -85.4444
7.1251 | -85.4456
7.1255 | -85.4472
7.1258 | -85.4488
7.1261 | -85.4502
7.1264 | -85.4508
7.1267 | -85.4498
7.1271 | -85.4504
7.1274 | -85.4516
7.1277 | -85.4532
7.128 | -85.4547
7.1284 | -85.456
7.1287 | -85.4568
7.129 | -85.4576
7.1293 | -85.4588
7.1296 | -85.4601
7.13 | -85.461
7.1303 | -85.462
7.1306 | -85.462
7.1309 | -85.4631
7.1312 | -85.464
7.1316 | -85.4651
7.1319 | -85.4658
7.1322 | -85.4653
7.1325 | -85.4662
7.1329 | -85.4678
7.1332 | -85.4698
7.1335 | -85.4707
7.1338 | -85.4714
7.1341 | -85.4728
7.1345 | -85.4744
7.1348 | -85.4752
7.1351 | -85.4756
7.1354 | -85.4766
7.1357 | -85.4785
7.1361 | -85.4804
7.1364 | -85.481
7.1367 | -85.483
7.137 | -85.4851
7.1373 | -85.4862
7.1377 | -85.4874
7.138 | -85.4882
7.1383 | -85.4887
7.1386 | -85.489
7.139 | -85.4898
7.1393 | -85.4905
7.1396 | -85.4915
7.1399 | -85.4929
7.1402 | -85.4941
7.1406 | -85.4948
7.1409 | -85.4961
7.1412 | -85.498
7.1415 | -85.4994
7.1418 | -85.5001
7.1422 | -85.501
7.1425 | -85.502
7.1428 | -85.503
7.1431 | -85.5037
7.1435 | -85.5047
7.1438 | -85.506
7.1441 | -85.5068
7.1444 | -85.5075
7.1447 | -85.5083
7.1451 | -85.5092
7.1454 | -85.5106
7.1457 | -85.5119
7.146 | -85.5131
7.1463 | -85.5141
7.1467 | -85.5154
7.147 | -85.5166
7.1473 | -85.5178
7.1476 | -85.5191
7.148 | -85.5185
7.1483 | -85.5193
7.1486 | -85.5203
7.1489 | -85.5214
7.1492 | -85.5224
7.1496 | -85.5234
7.1499 | -85.5246
7.1502 | -85.5256
7.1505 | -85.5265
7.1508 | -85.5274
7.1512 | -85.5283
7.1515 | -85.5291
7.1518 | -85.5298
7.1521 | -85.5305
7.1525 | -85.5313
7.1528 | -85.5322
7.1531 | -85.5324
7.1534 | -85.5323
7.1537 | -85.5334
7.1541 | -85.5344
7.1544 | -85.5356
7.1547 | -85.5367
7.155 | -85.5378
7.1553 | -85.5401
7.1557 | -85.5424
7.156 | -85.5439
7.1563 | -85.545
7.1566 | -85.546
7.157 | -85.5474
7.1573 | -85.5471
7.1576 | -85.547
7.1579 | -85.548
7.1582 | -85.5494
7.1586 | -85.5509
7.1589 | -85.5523
7.1592 | -85.553
7.1595 | -85.5535
7.1598 | -85.5543
7.1602 | -85.5552
7.1605 | -85.5565
7.1608 | -85.557
7.1611 | -85.5572
7.1614 | -85.5577
7.1618 | -85.5583
7.1621 | -85.559
7.1624 | -85.5599
7.1627 | -85.5601
7.1631 | -85.5595
7.1634 | -85.5598
7.1637 | -85.5605
7.164 | -85.5615
7.1643 | -85.5624
7.1647 | -85.5632
7.165 | -85.5643
7.1653 | -85.565
7.1656 | -85.5658
7.1659 | -85.5668
7.1663 | -85.5677
7.1666 | -85.5688
7.1669 | -85.5703
7.1672 | -85.5708
7.1676 | -85.5718
7.1679 | -85.5729
7.1682 | -85.5738
7.1685 | -85.5744
7.1688 | -85.5754
7.1692 | -85.5761
7.1695 | -85.5767
7.1698 | -85.5775
7.1701 | -85.5794
7.1704 | -85.5816
7.1708 | -85.5827
7.1711 | -85.5834
7.1714 | -85.5841
7.1717 | -85.5846
7.1721 | -85.5848
7.1724 | -85.5851
7.1727 | -85.5855
7.173 | -85.5862
7.1733 | -85.5875
7.1737 | -85.5882
7.174 | -85.589
7.1743 | -85.5893
7.1746 | -85.5897
7.1749 | -85.5906
7.1753 | -85.5914
7.1756 | -85.592
7.1759 | -85.5929
7.1762 | -85.594
7.1766 | -85.595
7.1769 | -85.5955
7.1772 | -85.5961
7.1775 | -85.5962
7.1778 | -85.5958
7.1782 | -85.5959
7.1785 | -85.5973
7.1788 | -85.5994
7.1791 | -85.6002
7.1794 | -85.6005
7.1798 | -85.6009
7.1801 | -85.602
7.1804 | -85.6028
7.1807 | -85.6034
7.181 | -85.6031
7.1814 | -85.6028
7.1817 | -85.6036
7.182 | -85.604
7.1823 | -85.6051
7.1827 | -85.6058
7.183 | -85.6064
7.1833 | -85.6069
7.1836 | -85.6075
7.1839 | -85.6084
7.1843 | -85.6096
7.1846 | -85.6101
7.1849 | -85.6109
7.1852 | -85.6114
7.1855 | -85.6114
7.1859 | -85.6114
7.1862 | -85.6125
7.1865 | -85.613
7.1868 | -85.6138
7.1872 | -85.6145
7.1875 | -85.615
7.1878 | -85.6161
7.1881 | -85.6165
7.1884 | -85.617
7.1888 | -85.6178
7.1891 | -85.6184
7.1894 | -85.6188
7.1897 | -85.6197
7.19 | -85.6205
7.1904 | -85.6208
7.1907 | -85.6206
7.191 | -85.6213
7.1913 | -85.6229
7.1917 | -85.6244
7.192 | -85.6247
7.1923 | -85.6249
7.1926 | -85.6256
7.1929 | -85.626
7.1933 | -85.6276
7.1936 | -85.6288
7.1939 | -85.629
7.1942 | -85.6285
7.1945 | -85.6289
7.1949 | -85.6302
7.1952 | -85.6304
7.1955 | -85.6311
7.1958 | -85.6315
7.1962 | -85.6316
7.1965 | -85.632
7.1968 | -85.6327
7.1971 | -85.6337
7.1974 | -85.6344
7.1978 | -85.6349
7.1981 | -85.6361
7.1984 | -85.6364
7.1987 | -85.6361
7.199 | -85.6357
7.1994 | -85.636
7.1997 | -85.6368
7.2 | -85.6383
7.2003 | -85.6388
7.2007 | -85.6392
7.201 | -85.6395
7.2013 | -85.6393
7.2016 | -85.6382
7.2019 | -85.6381
7.2023 | -85.6385
7.2026 | -85.6391
7.2029 | -85.6387
7.2032 | -85.6375
7.2035 | -85.6375
7.2039 | -85.6377
7.2042 | -85.6384
7.2045 | -85.6391
7.2048 | -85.6395
7.2051 | -85.6396
7.2055 | -85.6401
7.2058 | -85.6408
7.2061 | -85.6427
7.2064 | -85.6437
7.2068 | -85.644
7.2071 | -85.6448
7.2074 | -85.6457
7.2077 | -85.6462
7.208 | -85.6466
7.2084 | -85.6468
7.2087 | -85.6469
7.209 | -85.6471
7.2093 | -85.648
7.2096 | -85.6487
7.21 | -85.6492
7.2103 | -85.6494
7.2106 | -85.6507
7.2109 | -85.652
7.2113 | -85.6519
7.2116 | -85.652
7.2119 | -85.6522
7.2122 | -85.6522
7.2125 | -85.6524
7.2129 | -85.6522
7.2132 | -85.6524
7.2135 | -85.6527
7.2138 | -85.6528
7.2141 | -85.6537
7.2145 | -85.6543
7.2148 | -85.6548
7.2151 | -85.6553
7.2154 | -85.6559
7.2158 | -85.6563
7.2161 | -85.6566
7.2164 | -85.6572
7.2167 | -85.6574
7.217 | -85.6579
7.2174 | -85.6587
7.2177 | -85.6591
7.218 | -85.6587
7.2183 | -85.6591
7.2186 | -85.6597
7.219 | -85.6603
7.2193 | -85.6607
7.2196 | -85.6611
7.2199 | -85.6611
7.2203 | -85.6613
7.2206 | -85.6612
7.2209 | -85.6615
7.2212 | -85.6621
7.2215 | -85.6621
7.2219 | -85.6618
7.2222 | -85.6614
7.2225 | -85.6617
7.2228 | -85.6616
7.2231 | -85.6615
7.2235 | -85.6614
7.2238 | -85.6611
7.2241 | -85.6607
7.2244 | -85.6599
7.2248 | -85.6598
7.2251 | -85.6604
7.2254 | -85.662
7.2257 | -85.6617
7.226 | -85.6615
7.2264 | -85.6622
7.2267 | -85.6641
7.227 | -85.6656
7.2273 | -85.6657
7.2276 | -85.6657
7.228 | -85.6658
7.2283 | -85.6658
7.2286 | -85.6656
7.2289 | -85.6657
7.2292 | -85.6659
7.2296 | -85.6661
7.2299 | -85.6664
7.2302 | -85.6663
7.2305 | -85.6657
7.2309 | -85.6655
7.2312 | -85.6653
7.2315 | -85.665
7.2318 | -85.6649
7.2321 | -85.6649
7.2325 | -85.6648
7.2328 | -85.6648
7.2331 | -85.665
7.2334 | -85.6651
7.2337 | -85.665
7.2341 | -85.6652
7.2344 | -85.6652
7.2347 | -85.6656
7.235 | -85.6659
7.2354 | -85.6665
7.2357 | -85.6666
7.236 | -85.6667
7.2363 | -85.667
7.2366 | -85.6669
7.237 | -85.6671
7.2373 | -85.6672
7.2376 | -85.6671
7.2379 | -85.6671
7.2382 | -85.6673
7.2386 | -85.6673
7.2389 | -85.6667
7.2392 | -85.6667
7.2395 | -85.667
7.2399 | -85.6674
7.2402 | -85.6678
7.2405 | -85.6679
7.2408 | -85.6679
7.2411 | -85.6678
7.2415 | -85.6679
7.2418 | -85.6683
7.2421 | -85.6689
7.2424 | -85.6687
7.2427 | -85.6692
7.2431 | -85.6698
7.2434 | -85.67
7.2437 | -85.6699
7.244 | -85.6694
7.2444 | -85.6696
7.2447 | -85.67
7.245 | -85.6701
7.2453 | -85.6705
7.2456 | -85.6707
7.246 | -85.6705
7.2463 | -85.6705
7.2466 | -85.6705
7.2469 | -85.6712
7.2472 | -85.6729
7.2476 | -85.6735
7.2479 | -85.6735
7.2482 | -85.6733
7.2485 | -85.6729
7.2489 | -85.6733
7.2492 | -85.6731
7.2495 | -85.673
7.2498 | -85.6733
7.2501 | -85.6732
7.2505 | -85.6727
7.2508 | -85.6723
7.2511 | -85.6709
7.2514 | -85.6706
7.2517 | -85.6697
7.2521 | -85.6697
7.2524 | -85.6699
7.2527 | -85.6694
7.253 | -85.6692
7.2533 | -85.6692
7.2537 | -85.6692
7.254 | -85.6691
7.2543 | -85.6697
7.2546 | -85.6707
7.255 | -85.6707
7.2553 | -85.6706
7.2556 | -85.6706
7.2559 | -85.6705
7.2562 | -85.6705
7.2566 | -85.6709
7.2569 | -85.6711
7.2572 | -85.6717
7.2575 | -85.6717
7.2578 | -85.6718
7.2582 | -85.6719
7.2585 | -85.6717
7.2588 | -85.6712
7.2591 | -85.6704
7.2595 | -85.6697
7.2598 | -85.6691
7.2601 | -85.669
7.2604 | -85.6689
7.2607 | -85.6688
7.2611 | -85.6689
7.2614 | -85.6688
7.2617 | -85.6686
7.262 | -85.6701
7.2623 | -85.6705
7.2627 | -85.6703
7.263 | -85.6702
7.2633 | -85.6707
7.2636 | -85.6709
7.264 | -85.6706
7.2643 | -85.6706
7.2646 | -85.6709
7.2649 | -85.6704
7.2652 | -85.6712
7.2656 | -85.6731
7.2659 | -85.6736
7.2662 | -85.6735
7.2665 | -85.6727
7.2668 | -85.6726
7.2672 | -85.6726
7.2675 | -85.6724
7.2678 | -85.6724
7.2681 | -85.6732
7.2685 | -85.6733
7.2688 | -85.6747
7.2691 | -85.6755
7.2694 | -85.6756
7.2697 | -85.6751
7.2701 | -85.6736
7.2704 | -85.6727
7.2707 | -85.6726
7.271 | -85.6723
7.2713 | -85.6721
7.2717 | -85.672
7.272 | -85.672
7.2723 | -85.6719
7.2726 | -85.6721
7.2729 | -85.6729
7.2733 | -85.6735
7.2736 | -85.6738
7.2739 | -85.6741
7.2742 | -85.675
7.2746 | -85.6746
7.2749 | -85.6744
7.2752 | -85.6742
7.2755 | -85.6742
7.2758 | -85.6741
7.2762 | -85.6742
7.2765 | -85.6752
7.2768 | -85.6754
7.2771 | -85.6753
7.2774 | -85.6759
7.2778 | -85.6761
7.2781 | -85.6771
7.2784 | -85.6777
7.2787 | -85.6779
7.2791 | -85.6777
7.2794 | -85.6779
7.2797 | -85.6777
7.28 | -85.6777
7.2803 | -85.6777
7.2807 | -85.6776
7.281 | -85.6772
7.2813 | -85.6763
7.2816 | -85.6764
7.2819 | -85.6765
7.2823 | -85.6765
7.2826 | -85.6765
7.2829 | -85.676
7.2832 | -85.6753
7.2836 | -85.6751
7.2839 | -85.676
7.2842 | -85.6766
7.2845 | -85.6771
7.2848 | -85.6775
7.2852 | -85.6783
7.2855 | -85.6784
7.2858 | -85.6784
7.2861 | -85.6782
7.2864 | -85.6783
7.2868 | -85.6784
7.2871 | -85.6784
7.2874 | -85.678
7.2877 | -85.6787
7.2881 | -85.6802
7.2884 | -85.6806
7.2887 | -85.6807
7.289 | -85.6808
7.2893 | -85.6804
7.2897 | -85.68
7.29 | -85.6795
7.2903 | -85.6787
7.2906 | -85.6784
7.2909 | -85.678
7.2913 | -85.6769
7.2916 | -85.6762
7.2919 | -85.676
7.2922 | -85.6757
7.2926 | -85.6745
7.2929 | -85.6746
7.2932 | -85.6755
7.2935 | -85.676
7.2938 | -85.6764
7.2942 | -85.6765
7.2945 | -85.6762
7.2948 | -85.6766
7.2951 | -85.6767
7.2954 | -85.6764
7.2958 | -85.6763
7.2961 | -85.6768
7.2964 | -85.6771
7.2967 | -85.6777
7.297 | -85.6783
7.2974 | -85.6793
7.2977 | -85.6796
7.298 | -85.6803
7.2983 | -85.6799
7.2987 | -85.6801
7.299 | -85.6795
7.2993 | -85.6796
7.2996 | -85.6801
7.2999 | -85.6802
7.3003 | -85.6802
7.3006 | -85.6803
7.3009 | -85.6802
7.3012 | -85.6802
7.3015 | -85.68
7.3019 | -85.68
7.3022 | -85.6792
7.3025 | -85.6795
7.3028 | -85.6795
7.3032 | -85.6796
7.3035 | -85.6796
7.3038 | -85.6796
7.3041 | -85.6796
7.3044 | -85.6795
7.3048 | -85.6796
7.3051 | -85.6794
7.3054 | -85.6798
7.3057 | -85.6794
7.306 | -85.6794
7.3064 | -85.6798
7.3067 | -85.6803
7.307 | -85.6802
7.3073 | -85.6809
7.3077 | -85.6814
7.308 | -85.6815
7.3083 | -85.6818
7.3086 | -85.682
7.3089 | -85.6822
7.3093 | -85.6827
7.3096 | -85.6827
7.3099 | -85.6824
7.3102 | -85.6833
7.3105 | -85.6842
7.3109 | -85.6841
7.3112 | -85.6844
7.3115 | -85.6849
7.3118 | -85.685
7.3122 | -85.6849
7.3125 | -85.6846
7.3128 | -85.6844
7.3131 | -85.6844
7.3134 | -85.6844
7.3138 | -85.6841
7.3141 | -85.6836
7.3144 | -85.6835
7.3147 | -85.6837
7.315 | -85.6833
7.3154 | -85.683
7.3157 | -85.6829
7.316 | -85.6824
7.3163 | -85.6825
7.3167 | -85.6828
7.317 | -85.6837
7.3173 | -85.6842
7.3176 | -85.6843
7.3179 | -85.6844
7.3183 | -85.6846
7.3186 | -85.685
7.3189 | -85.6848
7.3192 | -85.6845
7.3195 | -85.6839
7.3199 | -85.6839
7.3202 | -85.6838
7.3205 | -85.6827
7.3208 | -85.6821
7.3211 | -85.6822
7.3215 | -85.6821
7.3218 | -85.6823
7.3221 | -85.6827
7.3224 | -85.6831
7.3228 | -85.683
7.3231 | -85.6825
7.3234 | -85.6822
7.3237 | -85.6818
7.324 | -85.6822
7.3244 | -85.6822
7.3247 | -85.6823
7.325 | -85.6832
7.3253 | -85.6835
7.3256 | -85.6836
7.326 | -85.6835
7.3263 | -85.6833
7.3266 | -85.6832
7.3269 | -85.6834
7.3273 | -85.6829
7.3276 | -85.6831
7.3279 | -85.6837
7.3282 | -85.6838
7.3285 | -85.6839
7.3289 | -85.6838
7.3292 | -85.6839
7.3295 | -85.684
7.3298 | -85.6839
7.3301 | -85.6835
7.3305 | -85.6839
7.3308 | -85.684
7.3311 | -85.6844
7.3314 | -85.6854
7.3318 | -85.6854
7.3321 | -85.685
7.3324 | -85.6848
7.3327 | -85.685
7.333 | -85.6851
7.3334 | -85.686
7.3337 | -85.6861
7.334 | -85.6862
7.3343 | -85.6867
7.3346 | -85.6867
7.335 | -85.6867
7.3353 | -85.6875
7.3356 | -85.6875
7.3359 | -85.6871
7.3363 | -85.6867
7.3366 | -85.6867
7.3369 | -85.687
7.3372 | -85.6877
7.3375 | -85.6882
7.3379 | -85.6884
7.3382 | -85.6887
7.3385 | -85.6896
7.3388 | -85.6899
7.3391 | -85.6901
7.3395 | -85.6908
7.3398 | -85.6916
7.3401 | -85.6917
7.3404 | -85.6914
7.3407 | -85.6917
7.3411 | -85.6919
7.3414 | -85.6922
7.3417 | -85.6923
7.342 | -85.6925
7.3424 | -85.6923
7.3427 | -85.6931
7.343 | -85.6934
7.3433 | -85.6936
7.3436 | -85.6933
7.344 | -85.6931
7.3443 | -85.6933
7.3446 | -85.6931
7.3449 | -85.6933
7.3452 | -85.6938
7.3456 | -85.6943
7.3459 | -85.6946
7.3462 | -85.695
7.3465 | -85.695
7.3469 | -85.6952
7.3472 | -85.6954
7.3475 | -85.6958
7.3478 | -85.6957
7.3481 | -85.6967
7.3485 | -85.6972
7.3488 | -85.6974
7.3491 | -85.6976
7.3494 | -85.698
7.3497 | -85.6973
7.3501 | -85.6966
7.3504 | -85.6971
7.3507 | -85.6974
7.351 | -85.6978
7.3514 | -85.6985
7.3517 | -85.6992
7.352 | -85.6998
7.3523 | -85.7004
7.3526 | -85.7015
7.353 | -85.702
7.3533 | -85.7021
7.3536 | -85.7021
7.3539 | -85.7022
7.3542 | -85.703
7.3546 | -85.7037
7.3549 | -85.7042
7.3552 | -85.7047
7.3555 | -85.7048
7.3559 | -85.7051
7.3562 | -85.7055
7.3565 | -85.7054
7.3568 | -85.7054
7.3571 | -85.7057
7.3575 | -85.7058
7.3578 | -85.7062
7.3581 | -85.7068
7.3584 | -85.7073
7.3587 | -85.7083
7.3591 | -85.7085
7.3594 | -85.7098
7.3597 | -85.7109
7.36 | -85.712
7.3604 | -85.7128
7.3607 | -85.7131
7.361 | -85.7137
7.3613 | -85.714
7.3616 | -85.7148
7.362 | -85.7156
7.3623 | -85.716
7.3626 | -85.7165
7.3629 | -85.7171
7.3632 | -85.7174
7.3636 | -85.718
7.3639 | -85.7187
7.3642 | -85.7192
7.3645 | -85.7191
7.3648 | -85.7195
7.3652 | -85.7201
7.3655 | -85.7204
7.3658 | -85.721
7.3661 | -85.722
7.3665 | -85.7224
7.3668 | -85.7229
7.3671 | -85.7243
7.3674 | -85.7255
7.3677 | -85.726
7.3681 | -85.7262
7.3684 | -85.7269
7.3687 | -85.7275
7.369 | -85.7284
7.3693 | -85.7292
7.3697 | -85.7297
7.37 | -85.7304
7.3703 | -85.7307
7.3706 | -85.7313
7.371 | -85.7326
7.3713 | -85.7333
7.3716 | -85.7339
7.3719 | -85.7347
7.3722 | -85.7355
7.3726 | -85.7358
7.3729 | -85.7366
7.3732 | -85.7383
7.3735 | -85.7391
7.3738 | -85.7396
7.3742 | -85.7404
7.3745 | -85.741
7.3748 | -85.7422
7.3751 | -85.7443
7.3755 | -85.7453
7.3758 | -85.7457
7.3761 | -85.7464
7.3764 | -85.7477
7.3767 | -85.7485
7.3771 | -85.7492
7.3774 | -85.7502
7.3777 | -85.7509
7.378 | -85.7517
7.3783 | -85.7522
7.3787 | -85.7529
7.379 | -85.7534
7.3793 | -85.754
7.3796 | -85.7556
7.38 | -85.7565
7.3803 | -85.7578
7.3806 | -85.76
7.3809 | -85.7612
7.3812 | -85.762
7.3816 | -85.7629
7.3819 | -85.7633
7.3822 | -85.764
7.3825 | -85.7647
7.3828 | -85.7653
7.3832 | -85.766
7.3835 | -85.7665
7.3838 | -85.767
7.3841 | -85.7679
7.3845 | -85.769
7.3848 | -85.7699
7.3851 | -85.7714
7.3854 | -85.7725
7.3857 | -85.7722
7.3861 | -85.7722
7.3864 | -85.7729
7.3867 | -85.7734
7.387 | -85.7739
7.3873 | -85.775
7.3877 | -85.7765
7.388 | -85.7768
7.3883 | -85.7771
7.3886 | -85.778
7.3889 | -85.7792
7.3893 | -85.7803
7.3896 | -85.7811
7.3899 | -85.7819
7.3902 | -85.7827
7.3906 | -85.7838
7.3909 | -85.7844
7.3912 | -85.7853
7.3915 | -85.7861
7.3918 | -85.7865
7.3922 | -85.787
7.3925 | -85.7884
7.3928 | -85.7891
7.3931 | -85.7902
7.3934 | -85.791
7.3938 | -85.7918
7.3941 | -85.7926
7.3944 | -85.7934
7.3947 | -85.7936
7.3951 | -85.7936
7.3954 | -85.794
7.3957 | -85.7949
7.396 | -85.7952
7.3963 | -85.7955
7.3967 | -85.7962
7.397 | -85.797
7.3973 | -85.7976
7.3976 | -85.7979
7.3979 | -85.799
7.3983 | -85.8006
7.3986 | -85.8015
7.3989 | -85.8024
7.3992 | -85.8035
7.3996 | -85.8043
7.3999 | -85.8055
7.4002 | -85.8063
7.4005 | -85.8071
7.4008 | -85.8082
7.4012 | -85.8094
7.4015 | -85.8102
7.4018 | -85.8113
7.4021 | -85.8125
7.4024 | -85.8135
7.4028 | -85.8144
7.4031 | -85.815
7.4034 | -85.816
7.4037 | -85.8181
7.4041 | -85.8193
7.4044 | -85.8203
7.4047 | -85.8212
7.405 | -85.8224
7.4053 | -85.8238
7.4057 | -85.825
7.406 | -85.8269
7.4063 | -85.8288
7.4066 | -85.8301
7.4069 | -85.8311
7.4073 | -85.832
7.4076 | -85.8339
7.4079 | -85.8356
7.4082 | -85.8368
7.4085 | -85.8374
7.4089 | -85.8378
7.4092 | -85.8383
7.4095 | -85.8398
7.4098 | -85.8431
7.4102 | -85.8446
7.4105 | -85.8459
7.4108 | -85.8472
7.4111 | -85.848
7.4114 | -85.8487
7.4118 | -85.8494
7.4121 | -85.8497
7.4124 | -85.8498
7.4127 | -85.8509
7.413 | -85.852
7.4134 | -85.8525
7.4137 | -85.8525
7.414 | -85.8532
7.4143 | -85.8543
7.4147 | -85.8554
7.415 | -85.8568
7.4153 | -85.8582
7.4156 | -85.8589
7.4159 | -85.86
7.4163 | -85.8614
7.4166 | -85.863
7.4169 | -85.8641
7.4172 | -85.8652
7.4175 | -85.8662
7.4179 | -85.8674
7.4182 | -85.8685
7.4185 | -85.8695
7.4188 | -85.8703
7.4192 | -85.8711
7.4195 | -85.8721
7.4198 | -85.8732
7.4201 | -85.8744
7.4204 | -85.8754
7.4208 | -85.8764
7.4211 | -85.8777
7.4214 | -85.8789
7.4217 | -85.8797
7.422 | -85.8806
7.4224 | -85.8814
7.4227 | -85.8822
7.423 | -85.8834
7.4233 | -85.8846
7.4237 | -85.8857
7.424 | -85.8868
7.4243 | -85.8882
7.4246 | -85.8894
7.4249 | -85.8904
7.4253 | -85.8916
7.4256 | -85.8926
7.4259 | -85.8944
7.4262 | -85.8951
7.4265 | -85.8964
7.4269 | -85.8977
7.4272 | -85.8988
7.4275 | -85.9001
7.4278 | -85.9016
7.4282 | -85.9027
7.4285 | -85.9039
7.4288 | -85.9051
7.4291 | -85.9056
7.4294 | -85.906
7.4298 | -85.907
7.4301 | -85.9081
7.4304 | -85.9094
7.4307 | -85.9106
7.431 | -85.9119
7.4314 | -85.9129
7.4317 | -85.9143
7.432 | -85.9155
7.4323 | -85.9169
7.4326 | -85.9181
7.433 | -85.9196
7.4333 | -85.9216
7.4336 | -85.9229
7.4339 | -85.9241
7.4343 | -85.9253
7.4346 | -85.9262
7.4349 | -85.9271
7.4352 | -85.929
7.4355 | -85.9302
7.4359 | -85.9317
7.4362 | -85.933
7.4365 | -85.9339
7.4368 | -85.935
7.4371 | -85.9368
7.4375 | -85.9389
7.4378 | -85.9403
7.4381 | -85.9413
7.4384 | -85.9427
7.4388 | -85.9439
7.4391 | -85.9444
7.4394 | -85.9452
7.4397 | -85.9464
7.44 | -85.9475
7.4404 | -85.9488
7.4407 | -85.9499
7.441 | -85.9511
7.4413 | -85.9522
7.4416 | -85.954
7.442 | -85.9556
7.4423 | -85.9571
7.4426 | -85.9584
7.4429 | -85.96
7.4433 | -85.9613
7.4436 | -85.9628
7.4439 | -85.9651
7.4442 | -85.9668
7.4445 | -85.9692
7.4449 | -85.9705
7.4452 | -85.972
7.4455 | -85.9734
7.4458 | -85.9747
7.4461 | -85.9761
7.4465 | -85.9776
7.4468 | -85.9787
7.4471 | -85.9804
7.4474 | -85.9821
7.4478 | -85.9834
7.4481 | -85.9849
7.4484 | -85.9862
7.4487 | -85.9876
7.449 | -85.989
7.4494 | -85.9901
7.4497 | -85.9914
7.45 | -85.9928
7.4503 | -85.9941
7.4506 | -85.9955
7.451 | -85.9972
7.4513 | -85.9986
7.4516 | -86
7.4519 | -86.0034
7.4523 | -86.0072
7.4526 | -86.0089
7.4529 | -86.0109
7.4532 | -86.0122
7.4535 | -86.013
7.4539 | -86.015
7.4542 | -86.0166
7.4545 | -86.0181
7.4548 | -86.0196
7.4551 | -86.0215
7.4555 | -86.0225
7.4558 | -86.0238
7.4561 | -86.025
7.4564 | -86.0266
7.4567 | -86.0282
7.4571 | -86.0298
7.4574 | -86.0317
7.4577 | -86.0337
7.458 | -86.0355
7.4584 | -86.0368
7.4587 | -86.0376
7.459 | -86.0386
7.4593 | -86.0398
7.4596 | -86.0413
7.46 | -86.043
7.4603 | -86.0447
7.4606 | -86.0473
7.4609 | -86.0497
7.4612 | -86.051
7.4616 | -86.0526
7.4619 | -86.0544
7.4622 | -86.0559
7.4625 | -86.0573
7.4629 | -86.0586
7.4632 | -86.061
7.4635 | -86.0625
7.4638 | -86.0643
7.4641 | -86.0668
7.4645 | -86.0684
7.4648 | -86.0704
7.4651 | -86.0726
7.4654 | -86.0752
7.4657 | -86.077
7.4661 | -86.0785
7.4664 | -86.0805
7.4667 | -86.0822
7.467 | -86.084
7.4674 | -86.0858
7.4677 | -86.0874
7.468 | -86.0892
7.4683 | -86.0909
7.4686 | -86.0924
7.469 | -86.0941
7.4693 | -86.0955
7.4696 | -86.0974
7.4699 | -86.0992
7.4702 | -86.1008
7.4706 | -86.1025
7.4709 | -86.1041
7.4712 | -86.1056
7.4715 | -86.1074
7.4719 | -86.1087
7.4722 | -86.1098
7.4725 | -86.1118
7.4728 | -86.1134
7.4731 | -86.1151
7.4735 | -86.1164
7.4738 | -86.1179
7.4741 | -86.1198
7.4744 | -86.1217
7.4747 | -86.123
7.4751 | -86.1245
7.4754 | -86.1261
7.4757 | -86.1277
7.476 | -86.1296
7.4764 | -86.1318
7.4767 | -86.1337
7.477 | -86.1358
7.4773 | -86.1379
7.4776 | -86.1394
7.478 | -86.141
7.4783 | -86.1436
7.4786 | -86.1462
7.4789 | -86.1481
7.4792 | -86.15
7.4796 | -86.1519
7.4799 | -86.1552
7.4802 | -86.1587
7.4805 | -86.1608
7.4808 | -86.1621
7.4812 | -86.1636
7.4815 | -86.1654
7.4818 | -86.1674
7.4821 | -86.1686
7.4825 | -86.1701
7.4828 | -86.1723
7.4831 | -86.1741
7.4834 | -86.1758
7.4837 | -86.1778
7.4841 | -86.1808
7.4844 | -86.1826
7.4847 | -86.1844
7.485 | -86.1862
7.4853 | -86.1882
7.4857 | -86.1899
7.486 | -86.1917
7.4863 | -86.1937
7.4866 | -86.1952
7.487 | -86.1965
7.4873 | -86.1985
7.4876 | -86.2007
7.4879 | -86.2024
7.4882 | -86.2043
7.4886 | -86.2065
7.4889 | -86.2083
7.4892 | -86.2101
7.4895 | -86.2116
7.4898 | -86.213
7.4902 | -86.2149
7.4905 | -86.2169
7.4908 | -86.2191
7.4911 | -86.2209
7.4915 | -86.2229
7.4918 | -86.2248
7.4921 | -86.2261
7.4924 | -86.2274
7.4927 | -86.2292
7.4931 | -86.2309
7.4934 | -86.2329
7.4937 | -86.2348
7.494 | -86.2365
7.4943 | -86.2382
7.4947 | -86.24
7.495 | -86.2416
7.4953 | -86.243
7.4956 | -86.245
7.496 | -86.2469
7.4963 | -86.2486
7.4966 | -86.251
7.4969 | -86.253
7.4972 | -86.2543
7.4976 | -86.2557
7.4979 | -86.2576
7.4982 | -86.2601
7.4985 | -86.262
7.4988 | -86.2644
7.4992 | -86.2661
7.4995 | -86.2683
7.4998 | -86.2699
7.5001 | -86.2721
7.5004 | -86.2741
7.5008 | -86.276
7.5011 | -86.2778
7.5014 | -86.2794
7.5017 | -86.2811
7.5021 | -86.2829
7.5024 | -86.2852
7.5027 | -86.2876
7.503 | -86.2903
7.5033 | -86.2925
7.5037 | -86.2942
7.504 | -86.2963
7.5043 | -86.2989
7.5046 | -86.3006
7.5049 | -86.3028
7.5053 | -86.3048
7.5056 | -86.3063
7.5059 | -86.308
7.5062 | -86.3098
7.5066 | -86.3116
7.5069 | -86.3138
7.5072 | -86.3168
7.5075 | -86.3202
7.5078 | -86.322
7.5082 | -86.3239
7.5085 | -86.3255
7.5088 | -86.3278
7.5091 | -86.3299
7.5094 | -86.3323
7.5098 | -86.3334
7.5101 | -86.3348
7.5104 | -86.3356
7.5107 | -86.3368
7.5111 | -86.3385
7.5114 | -86.3404
7.5117 | -86.3423
7.512 | -86.3438
7.5123 | -86.3459
7.5127 | -86.3484
7.513 | -86.3508
7.5133 | -86.3528
7.5136 | -86.355
7.5139 | -86.3572
7.5143 | -86.3593
7.5146 | -86.3613
7.5149 | -86.3632
7.5152 | -86.3646
7.5156 | -86.3677
7.5159 | -86.3698
7.5162 | -86.3726
7.5165 | -86.3744
7.5168 | -86.3762
7.5172 | -86.3781
7.5175 | -86.3805
7.5178 | -86.383
7.5181 | -86.3854
7.5184 | -86.3874
7.5188 | -86.3899
7.5191 | -86.3916
7.5194 | -86.3938
7.5197 | -86.3958
7.5201 | -86.3973
7.5204 | -86.3993
7.5207 | -86.4011
7.521 | -86.4028
7.5213 | -86.4051
7.5217 | -86.4067
7.522 | -86.4085
7.5223 | -86.4116
7.5226 | -86.4137
7.5229 | -86.4155
7.5233 | -86.4173
7.5236 | -86.4191
7.5239 | -86.4223
7.5242 | -86.4254
7.5245 | -86.4275
7.5249 | -86.429
7.5252 | -86.4305
7.5255 | -86.4322
7.5258 | -86.4343
7.5262 | -86.4354
7.5265 | -86.4365
7.5268 | -86.4381
7.5271 | -86.4396
7.5274 | -86.4409
7.5278 | -86.4429
7.5281 | -86.4451
7.5284 | -86.4467
7.5287 | -86.4487
7.529 | -86.4509
7.5294 | -86.4527
7.5297 | -86.4548
7.53 | -86.4569
7.5303 | -86.4587
7.5307 | -86.4608
7.531 | -86.4629
7.5313 | -86.4662
7.5316 | -86.4694
7.5319 | -86.4714
7.5323 | -86.4727
7.5326 | -86.4743
7.5329 | -86.4763
7.5332 | -86.4781
7.5335 | -86.4796
7.5339 | -86.4814
7.5342 | -86.4834
7.5345 | -86.4854
7.5348 | -86.4877
7.5352 | -86.4897
7.5355 | -86.4917
7.5358 | -86.4943
7.5361 | -86.4964
7.5364 | -86.4986
7.5368 | -86.5004
7.5371 | -86.5033
7.5374 | -86.5056
7.5377 | -86.5074
7.538 | -86.5095
7.5384 | -86.5115
7.5387 | -86.513
7.539 | -86.5149
7.5393 | -86.5168
7.5397 | -86.5188
7.54 | -86.5208
7.5403 | -86.5227
7.5406 | -86.5245
7.5409 | -86.5265
7.5413 | -86.5285
7.5416 | -86.5307
7.5419 | -86.5332
7.5422 | -86.5355
7.5425 | -86.5376
7.5429 | -86.5395
7.5432 | -86.5415
7.5435 | -86.544
7.5438 | -86.5463
7.5442 | -86.5486
7.5445 | -86.5508
7.5448 | -86.5532
7.5451 | -86.5553
7.5454 | -86.5575
7.5458 | -86.5596
7.5461 | -86.5614
7.5464 | -86.5628
7.5467 | -86.5651
7.547 | -86.5671
7.5474 | -86.5677
7.5477 | -86.5697
7.548 | -86.5722
7.5483 | -86.5743
7.5486 | -86.5764
7.549 | -86.5782
7.5493 | -86.58
7.5496 | -86.582
7.5499 | -86.5842
7.5503 | -86.5875
7.5506 | -86.59
7.5509 | -86.5918
7.5512 | -86.5943
7.5515 | -86.5971
7.5519 | -86.5992
7.5522 | -86.6008
7.5525 | -86.6039
7.5528 | -86.606
7.5531 | -86.6076
7.5535 | -86.6093
7.5538 | -86.6114
7.5541 | -86.6134
7.5544 | -86.6159
7.5548 | -86.6185
7.5551 | -86.6207
7.5554 | -86.6225
7.5557 | -86.6245
7.556 | -86.6263
7.5564 | -86.6286
7.5567 | -86.6311
7.557 | -86.6328
7.5573 | -86.6354
7.5576 | -86.6381
7.558 | -86.6413
7.5583 | -86.6437
7.5586 | -86.6458
7.5589 | -86.6474
7.5593 | -86.6486
7.5596 | -86.6503
7.5599 | -86.6531
7.5602 | -86.6556
7.5605 | -86.6571
7.5609 | -86.6591
7.5612 | -86.661
7.5615 | -86.6636
7.5618 | -86.6671
7.5621 | -86.6692
7.5625 | -86.6711
7.5628 | -86.6723
7.5631 | -86.6737
7.5634 | -86.6754
7.5638 | -86.678
7.5641 | -86.6807
7.5644 | -86.6824
7.5647 | -86.6842
7.565 | -86.6858
7.5654 | -86.6883
7.5657 | -86.6902
7.566 | -86.6922
7.5663 | -86.6943
7.5666 | -86.6962
7.567 | -86.6983
7.5673 | -86.7004
7.5676 | -86.7029
7.5679 | -86.7054
7.5682 | -86.709
7.5686 | -86.7119
7.5689 | -86.7141
7.5692 | -86.7161
7.5695 | -86.7175
7.5699 | -86.7193
7.5702 | -86.721
7.5705 | -86.7225
7.5708 | -86.724
7.5711 | -86.7259
7.5715 | -86.7279
7.5718 | -86.7298
7.5721 | -86.7333
7.5724 | -86.7354
7.5727 | -86.7373
7.5731 | -86.7394
7.5734 | -86.7417
7.5737 | -86.7436
7.574 | -86.7455
7.5744 | -86.7474
7.5747 | -86.749
7.575 | -86.7514
7.5753 | -86.7531
7.5756 | -86.7545
7.576 | -86.7565
7.5763 | -86.7582
7.5766 | -86.7599
7.5769 | -86.7617
7.5772 | -86.7643
7.5776 | -86.7662
7.5779 | -86.7676
7.5782 | -86.7688
7.5785 | -86.7708
7.5789 | -86.7736
7.5792 | -86.7751
7.5795 | -86.7764
7.5798 | -86.7779
7.5801 | -86.78
7.5805 | -86.782
7.5808 | -86.784
7.5811 | -86.7863
7.5814 | -86.7889
7.5817 | -86.7911
7.5821 | -86.7936
7.5824 | -86.796
7.5827 | -86.798
7.583 | -86.7998
7.5834 | -86.8016
7.5837 | -86.803
7.584 | -86.8045
7.5843 | -86.8062
7.5846 | -86.8083
7.585 | -86.8101
7.5853 | -86.8119
7.5856 | -86.8139
7.5859 | -86.8162
7.5862 | -86.8186
7.5866 | -86.8218
7.5869 | -86.8247
7.5872 | -86.8271
7.5875 | -86.8292
7.5879 | -86.8309
7.5882 | -86.8326
7.5885 | -86.8346
7.5888 | -86.8364
7.5891 | -86.8381
7.5895 | -86.8399
7.5898 | -86.8424
7.5901 | -86.8446
7.5904 | -86.8465
7.5907 | -86.8487
7.5911 | -86.851
7.5914 | -86.8558
7.5917 | -86.8589
7.592 | -86.8603
7.5923 | -86.8619
7.5927 | -86.8635
7.593 | -86.8654
7.5933 | -86.8671
7.5936 | -86.8689
7.594 | -86.8705
7.5943 | -86.8724
7.5946 | -86.8742
7.5949 | -86.8762
7.5952 | -86.8781
7.5956 | -86.8802
7.5959 | -86.8826
7.5962 | -86.8848
7.5965 | -86.8873
7.5968 | -86.8892
7.5972 | -86.8908
7.5975 | -86.8926
7.5978 | -86.8948
7.5981 | -86.8966
7.5985 | -86.8985
7.5988 | -86.9004
7.5991 | -86.9024
7.5994 | -86.9049
7.5997 | -86.9069
7.6001 | -86.9086
7.6004 | -86.9108
7.6007 | -86.9128
7.601 | -86.9154
7.6013 | -86.9185
7.6017 | -86.92
7.602 | -86.9219
7.6023 | -86.9238
7.6026 | -86.9254
7.603 | -86.9266
7.6033 | -86.9289
7.6036 | -86.9317
7.6039 | -86.9339
7.6042 | -86.9356
7.6046 | -86.9376
7.6049 | -86.9392
7.6052 | -86.9408
7.6055 | -86.9434
7.6058 | -86.9455
7.6062 | -86.9474
7.6065 | -86.9492
7.6068 | -86.9513
7.6071 | -86.9531
7.6075 | -86.9546
7.6078 | -86.9564
7.6081 | -86.9582
7.6084 | -86.9603
7.6087 | -86.9623
7.6091 | -86.9644
7.6094 | -86.9661
7.6097 | -86.9677
7.61 | -86.9691
7.6103 | -86.9716
7.6107 | -86.9741
7.611 | -86.9763
7.6113 | -86.9778
7.6116 | -86.9798
7.612 | -86.9816
7.6123 | -86.9833
7.6126 | -86.9848
7.6129 | -86.9866
7.6132 | -86.9887
7.6136 | -86.9905
7.6139 | -86.9922
7.6142 | -86.995
7.6145 | -86.9967
7.6148 | -86.9987
7.6152 | -87.0002
7.6155 | -87.0017
7.6158 | -87.0034
7.6161 | -87.0047
7.6164 | -87.0067
7.6168 | -87.0089
7.6171 | -87.0105
7.6174 | -87.013
7.6177 | -87.0152
7.6181 | -87.0166
7.6184 | -87.0185
7.6187 | -87.0204
7.619 | -87.0219
7.6193 | -87.023
7.6197 | -87.0243
7.62 | -87.026
7.6203 | -87.0273
7.6206 | -87.0281
7.6209 | -87.0297
7.6213 | -87.0317
7.6216 | -87.0333
7.6219 | -87.0346
7.6222 | -87.0367
7.6226 | -87.0386
7.6229 | -87.0402
7.6232 | -87.0421
7.6235 | -87.0446
7.6238 | -87.0469
7.6242 | -87.0497
7.6245 | -87.0519
7.6248 | -87.0535
7.6251 | -87.0552
7.6254 | -87.0567
7.6258 | -87.0582
7.6261 | -87.0598
7.6264 | -87.0611
7.6267 | -87.0619
7.6271 | -87.0636
7.6274 | -87.0652
7.6277 | -87.0682
7.628 | -87.0705
7.6283 | -87.0721
7.6287 | -87.0729
7.629 | -87.0739
7.6293 | -87.0754
7.6296 | -87.0768
7.6299 | -87.0777
7.6303 | -87.0789
7.6306 | -87.0803
7.6309 | -87.0818
7.6312 | -87.0832
7.6316 | -87.0848
7.6319 | -87.0872
7.6322 | -87.089
7.6325 | -87.0906
7.6328 | -87.0922
7.6332 | -87.0942
7.6335 | -87.0961
7.6338 | -87.0974
7.6341 | -87.0986
7.6344 | -87.0994
7.6348 | -87.1003
7.6351 | -87.1018
7.6354 | -87.1039
7.6357 | -87.1061
7.6361 | -87.1077
7.6364 | -87.1092
7.6367 | -87.1104
7.637 | -87.1116
7.6373 | -87.1131
7.6377 | -87.1147
7.638 | -87.1161
7.6383 | -87.118
7.6386 | -87.1211
7.6389 | -87.1254
7.6393 | -87.1266
7.6396 | -87.129
7.6399 | -87.1305
7.6402 | -87.1316
7.6405 | -87.133
7.6409 | -87.1344
7.6412 | -87.1367
7.6415 | -87.1387
7.6418 | -87.14
7.6422 | -87.1405
7.6425 | -87.1418
7.6428 | -87.1433
7.6431 | -87.1449
7.6434 | -87.1477
7.6438 | -87.1494
7.6441 | -87.1511
7.6444 | -87.1525
7.6447 | -87.1547
7.645 | -87.1581
7.6454 | -87.1597
7.6457 | -87.1611
7.646 | -87.1625
7.6463 | -87.1645
7.6467 | -87.1661
7.647 | -87.1675
7.6473 | -87.1688
7.6476 | -87.1703
7.6479 | -87.1726
7.6483 | -87.1739
7.6486 | -87.1751
7.6489 | -87.1764
7.6492 | -87.1778
7.6495 | -87.1798
7.6499 | -87.1817
7.6502 | -87.1834
7.6505 | -87.1847
7.6508 | -87.1852
7.6512 | -87.187
7.6515 | -87.1892
7.6518 | -87.1909
7.6521 | -87.1932
7.6524 | -87.1952
7.6528 | -87.1969
7.6531 | -87.1982
7.6534 | -87.1996
7.6537 | -87.2011
7.654 | -87.2023
7.6544 | -87.2036
7.6547 | -87.205
7.655 | -87.2063
7.6553 | -87.2081
7.6557 | -87.2098
7.656 | -87.2107
7.6563 | -87.2118
7.6566 | -87.2136
7.6569 | -87.215
7.6573 | -87.2161
7.6576 | -87.2179
7.6579 | -87.2192
7.6582 | -87.2208
7.6585 | -87.2224
7.6589 | -87.2237
7.6592 | -87.2245
7.6595 | -87.2261
7.6598 | -87.2274
7.6601 | -87.2281
7.6605 | -87.2288
7.6608 | -87.2302
7.6611 | -87.2314
7.6614 | -87.2332
7.6618 | -87.2349
7.6621 | -87.2365
7.6624 | -87.2372
7.6627 | -87.2385
7.663 | -87.2401
7.6634 | -87.2413
7.6637 | -87.2424
7.664 | -87.244
7.6643 | -87.2453
7.6646 | -87.2467
7.665 | -87.2483
7.6653 | -87.2497
7.6656 | -87.2514
7.6659 | -87.2531
7.6663 | -87.255
7.6666 | -87.2563
7.6669 | -87.2575
7.6672 | -87.2588
7.6675 | -87.2611
7.6679 | -87.263
7.6682 | -87.2628
7.6685 | -87.2634
7.6688 | -87.2649
7.6691 | -87.2661
7.6695 | -87.2677
7.6698 | -87.2689
7.6701 | -87.2703
7.6704 | -87.2716
7.6708 | -87.2728
7.6711 | -87.2741
7.6714 | -87.2752
7.6717 | -87.2773
7.672 | -87.2795
7.6724 | -87.2819
7.6727 | -87.2837
7.673 | -87.2846
7.6733 | -87.2853
7.6736 | -87.2864
7.674 | -87.2882
7.6743 | -87.2895
7.6746 | -87.2902
7.6749 | -87.2918
7.6753 | -87.2931
7.6756 | -87.2954
7.6759 | -87.2975
7.6762 | -87.2981
7.6765 | -87.2994
7.6769 | -87.3004
7.6772 | -87.3016
7.6775 | -87.3026
7.6778 | -87.3047
7.6781 | -87.3078
7.6785 | -87.3097
7.6788 | -87.3108
7.6791 | -87.312
7.6794 | -87.3132
7.6798 | -87.3144
7.6801 | -87.3155
7.6804 | -87.3172
7.6807 | -87.3185
7.681 | -87.3193
7.6814 | -87.3205
7.6817 | -87.3215
7.682 | -87.3228
7.6823 | -87.3237
7.6826 | -87.3252
7.683 | -87.3265
7.6833 | -87.3279
7.6836 | -87.3291
7.6839 | -87.33
7.6842 | -87.3306
7.6846 | -87.3316
7.6849 | -87.3332
7.6852 | -87.335
7.6855 | -87.3358
7.6859 | -87.3374
7.6862 | -87.3389
7.6865 | -87.3396
7.6868 | -87.3404
7.6871 | -87.3407
7.6875 | -87.3416
7.6878 | -87.3435
7.6881 | -87.3442
7.6884 | -87.345
7.6887 | -87.3459
7.6891 | -87.3467
7.6894 | -87.3478
7.6897 | -87.3484
7.69 | -87.3497
7.6904 | -87.3513
7.6907 | -87.3522
7.691 | -87.3529
7.6913 | -87.3539
7.6916 | -87.3549
7.692 | -87.3558
7.6923 | -87.3561
7.6926 | -87.3556
7.6929 | -87.3561
7.6932 | -87.3573
7.6936 | -87.3589
7.6939 | -87.3615
7.6942 | -87.3634
7.6945 | -87.3643
7.6949 | -87.3652
7.6952 | -87.3663
7.6955 | -87.3675
7.6958 | -87.3683
7.6961 | -87.3692
7.6965 | -87.3699
7.6968 | -87.3707
7.6971 | -87.3716
7.6974 | -87.3737
7.6977 | -87.376
7.6981 | -87.3777
7.6984 | -87.3787
7.6987 | -87.3785
7.699 | -87.3795
7.6994 | -87.3804
7.6997 | -87.3815
7.7 | -87.3824
7.7003 | -87.3834
7.7006 | -87.3844
7.701 | -87.3859
7.7013 | -87.3867
7.7016 | -87.3875
7.7019 | -87.3885
7.7022 | -87.3904
7.7026 | -87.3922
7.7029 | -87.3935
7.7032 | -87.3946
7.7035 | -87.3954
7.7039 | -87.3956
7.7042 | -87.3951
7.7045 | -87.3966
7.7048 | -87.3977
7.7051 | -87.399
7.7055 | -87.4004
7.7058 | -87.4021
7.7061 | -87.4037
7.7064 | -87.4048
7.7067 | -87.4059
7.7071 | -87.4069
7.7074 | -87.409
7.7077 | -87.4112
7.708 | -87.4117
7.7083 | -87.4122
7.7087 | -87.4137
7.709 | -87.4159
7.7093 | -87.4168
7.7096 | -87.4177
7.71 | -87.4181
7.7103 | -87.4182
7.7106 | -87.4185
7.7109 | -87.4197
7.7112 | -87.4212
7.7116 | -87.4222
7.7119 | -87.4233
7.7122 | -87.4244
7.7125 | -87.4257
7.7128 | -87.4273
7.7132 | -87.4291
7.7135 | -87.4291
7.7138 | -87.43
7.7141 | -87.4308
7.7145 | -87.4313
7.7148 | -87.4316
7.7151 | -87.4326
7.7154 | -87.4334
7.7157 | -87.4344
7.7161 | -87.4359
7.7164 | -87.4383
7.7167 | -87.4404
7.717 | -87.4409
7.7173 | -87.4414
7.7177 | -87.4419
7.718 | -87.4432
7.7183 | -87.4441
7.7186 | -87.4445
7.719 | -87.4462
7.7193 | -87.4488
7.7196 | -87.4503
7.7199 | -87.4514
7.7202 | -87.4528
7.7206 | -87.4549
7.7209 | -87.4581
7.7212 | -87.459
7.7215 | -87.4594
7.7218 | -87.4601
7.7222 | -87.4607
7.7225 | -87.4613
7.7228 | -87.462
7.7231 | -87.4628
7.7235 | -87.4633
7.7238 | -87.464
7.7241 | -87.4654
7.7244 | -87.466
7.7247 | -87.4666
7.7251 | -87.4674
7.7254 | -87.4677
7.7257 | -87.4684
7.726 | -87.4688
7.7263 | -87.4695
7.7267 | -87.4704
7.727 | -87.4715
7.7273 | -87.4729
7.7276 | -87.4734
7.7279 | -87.474
7.7283 | -87.4744
7.7286 | -87.4751
7.7289 | -87.4758
7.7292 | -87.477
7.7296 | -87.4791
7.7299 | -87.4803
7.7302 | -87.4807
7.7305 | -87.4812
7.7308 | -87.4819
7.7312 | -87.4827
7.7315 | -87.4838
7.7318 | -87.4832
7.7321 | -87.4834
7.7324 | -87.4839
7.7328 | -87.4843
7.7331 | -87.4848
7.7334 | -87.4854
7.7337 | -87.4859
7.7341 | -87.487
7.7344 | -87.4872
7.7347 | -87.4882
7.735 | -87.4894
7.7353 | -87.49
7.7357 | -87.4899
7.736 | -87.4908
7.7363 | -87.4919
7.7366 | -87.4923
7.7369 | -87.4942
7.7373 | -87.4953
7.7376 | -87.4962
7.7379 | -87.4969
7.7382 | -87.4975
7.7386 | -87.4977
7.7389 | -87.4981
7.7392 | -87.4991
7.7395 | -87.5005
7.7398 | -87.5011
7.7402 | -87.5019
7.7405 | -87.5031
7.7408 | -87.5049
7.7411 | -87.5058
7.7414 | -87.5073
7.7418 | -87.5079
7.7421 | -87.5084
7.7424 | -87.5096
7.7427 | -87.5108
7.7431 | -87.5116
7.7434 | -87.5131
7.7437 | -87.5144
7.744 | -87.5163
7.7443 | -87.5178
7.7447 | -87.5187
7.745 | -87.5197
7.7453 | -87.5213
7.7456 | -87.5222
7.7459 | -87.5226
7.7463 | -87.5232
7.7466 | -87.5239
7.7469 | -87.5249
7.7472 | -87.5256
7.7476 | -87.5264
7.7479 | -87.5274
7.7482 | -87.5286
7.7485 | -87.5293
7.7488 | -87.5301
7.7492 | -87.5307
7.7495 | -87.5317
7.7498 | -87.5327
7.7501 | -87.5332
7.7504 | -87.534
7.7508 | -87.5354
7.7511 | -87.5362
7.7514 | -87.5367
7.7517 | -87.5376
7.752 | -87.5378
7.7524 | -87.5383
7.7527 | -87.5382
7.753 | -87.5384
7.7533 | -87.5389
7.7537 | -87.5396
7.754 | -87.5407
7.7543 | -87.5414
7.7546 | -87.5421
7.7549 | -87.5425
7.7553 | -87.5427
7.7556 | -87.5426
7.7559 | -87.5428
7.7562 | -87.5431
7.7565 | -87.5432
7.7569 | -87.5436
7.7572 | -87.544
7.7575 | -87.5444
7.7578 | -87.5452
7.7582 | -87.5456
7.7585 | -87.5459
7.7588 | -87.5472
7.7591 | -87.5484
7.7594 | -87.5491
7.7598 | -87.5478
7.7601 | -87.5485
7.7604 | -87.5495
7.7607 | -87.5502
7.761 | -87.5509
7.7614 | -87.5513
7.7617 | -87.552
7.762 | -87.5539
7.7623 | -87.5546
7.7627 | -87.555
7.763 | -87.5558
7.7633 | -87.5566
7.7636 | -87.5571
7.7639 | -87.5577
7.7643 | -87.5581
7.7646 | -87.5585
7.7649 | -87.5597
7.7652 | -87.5603
7.7655 | -87.5607
7.7659 | -87.5616
7.7662 | -87.5625
7.7665 | -87.5634
7.7668 | -87.5641
7.7672 | -87.5646
7.7675 | -87.565
7.7678 | -87.5653
7.7681 | -87.5659
7.7684 | -87.5667
7.7688 | -87.5674
7.7691 | -87.5675
7.7694 | -87.5682
7.7697 | -87.569
7.77 | -87.5687
7.7704 | -87.5691
7.7707 | -87.57
7.771 | -87.5705
7.7713 | -87.571
7.7717 | -87.5718
7.772 | -87.5722
7.7723 | -87.5724
7.7726 | -87.5742
7.7729 | -87.5748
7.7733 | -87.5765
7.7736 | -87.5773
7.7739 | -87.5777
7.7742 | -87.5787
7.7745 | -87.5788
7.7749 | -87.5793
7.7752 | -87.5795
7.7755 | -87.5801
7.7758 | -87.5805
7.7762 | -87.5816
7.7765 | -87.5824
7.7768 | -87.5829
7.7771 | -87.583
7.7774 | -87.5832
7.7778 | -87.5837
7.7781 | -87.5841
7.7784 | -87.5849
7.7787 | -87.5855
7.779 | -87.5855
7.7794 | -87.5856
7.7797 | -87.586
7.78 | -87.5864
7.7803 | -87.5871
7.7806 | -87.5882
7.781 | -87.5889
7.7813 | -87.5907
7.7816 | -87.5914
7.7819 | -87.5921
7.7823 | -87.5929
7.7826 | -87.5933
7.7829 | -87.5937
7.7832 | -87.5939
7.7835 | -87.5946
7.7839 | -87.5956
7.7842 | -87.5961
7.7845 | -87.5964
7.7848 | -87.5968
7.7851 | -87.597
7.7855 | -87.5971
7.7858 | -87.5975
7.7861 | -87.598
7.7864 | -87.5984
7.7868 | -87.5989
7.7871 | -87.5992
7.7874 | -87.5988
7.7877 | -87.5984
7.788 | -87.599
7.7884 | -87.5991
7.7887 | -87.5995
7.789 | -87.6008
7.7893 | -87.6011
7.7896 | -87.6016
7.79 | -87.6028
7.7903 | -87.6041
7.7906 | -87.6048
7.7909 | -87.6051
7.7913 | -87.6054
7.7916 | -87.6062
7.7919 | -87.6066
7.7922 | -87.6071
7.7925 | -87.6081
7.7929 | -87.6104
7.7932 | -87.6116
7.7935 | -87.611
7.7938 | -87.6102
7.7941 | -87.6103
7.7945 | -87.6106
7.7948 | -87.6108
7.7951 | -87.6113
7.7954 | -87.6117
7.7958 | -87.6111
7.7961 | -87.6113
7.7964 | -87.6119
7.7967 | -87.6122
7.797 | -87.6126
7.7974 | -87.6133
7.7977 | -87.614
7.798 | -87.6141
7.7983 | -87.6145
7.7986 | -87.6147
7.799 | -87.6144
7.7993 | -87.614
7.7996 | -87.6138
7.7999 | -87.6144
7.8002 | -87.6153
7.8006 | -87.6158
7.8009 | -87.6164
7.8012 | -87.6166
7.8015 | -87.6167
7.8019 | -87.6167
7.8022 | -87.617
7.8025 | -87.6173
7.8028 | -87.6177
7.8031 | -87.618
7.8035 | -87.6176
7.8038 | -87.6173
7.8041 | -87.6179
7.8044 | -87.619
7.8047 | -87.62
7.8051 | -87.6203
7.8054 | -87.6208
7.8057 | -87.621
7.806 | -87.6215
7.8064 | -87.6212
7.8067 | -87.6204
7.807 | -87.6212
7.8073 | -87.6221
7.8076 | -87.6225
7.808 | -87.6226
7.8083 | -87.6229
7.8086 | -87.6235
7.8089 | -87.6238
7.8092 | -87.6241
7.8096 | -87.6243
7.8099 | -87.6239
7.8102 | -87.6242
7.8105 | -87.6247
7.8109 | -87.6249
7.8112 | -87.6257
7.8115 | -87.6262
7.8118 | -87.6266
7.8121 | -87.6263
7.8125 | -87.6264
7.8128 | -87.6265
7.8131 | -87.6268
7.8134 | -87.6279
7.8137 | -87.6282
7.8141 | -87.6291
7.8144 | -87.6304
7.8147 | -87.6311
7.815 | -87.6303
7.8154 | -87.6297
7.8157 | -87.6313
7.816 | -87.6323
7.8163 | -87.6325
7.8166 | -87.6332
7.817 | -87.6345
7.8173 | -87.6351
7.8176 | -87.6352
7.8179 | -87.6357
7.8182 | -87.6365
7.8186 | -87.6368
7.8189 | -87.6376
7.8192 | -87.6381
7.8195 | -87.6382
7.8198 | -87.6385
7.8202 | -87.639
7.8205 | -87.6392
7.8208 | -87.6395
7.8211 | -87.6395
7.8215 | -87.6395
7.8218 | -87.6399
7.8221 | -87.64
7.8224 | -87.6408
7.8227 | -87.6422
7.8231 | -87.6441
7.8234 | -87.6452
7.8237 | -87.6463
7.824 | -87.6468
7.8243 | -87.647
7.8247 | -87.6475
7.825 | -87.6477
7.8253 | -87.6479
7.8256 | -87.6483
7.826 | -87.6486
7.8263 | -87.6489
7.8266 | -87.6491
7.8269 | -87.6495
7.8272 | -87.6506
7.8276 | -87.6512
7.8279 | -87.6513
7.8282 | -87.6513
7.8285 | -87.651
7.8288 | -87.6513
7.8292 | -87.6514
7.8295 | -87.6516
7.8298 | -87.6518
7.8301 | -87.6522
7.8305 | -87.6526
7.8308 | -87.6531
7.8311 | -87.6538
7.8314 | -87.6543
7.8317 | -87.6548
7.8321 | -87.6547
7.8324 | -87.6551
7.8327 | -87.655
7.833 | -87.6547
7.8333 | -87.6549
7.8337 | -87.6549
7.834 | -87.6552
7.8343 | -87.6552
7.8346 | -87.6553
7.835 | -87.6558
7.8353 | -87.6571
7.8356 | -87.6587
7.8359 | -87.6592
7.8362 | -87.6592
7.8366 | -87.6593
7.8369 | -87.6588
7.8372 | -87.6583
7.8375 | -87.6586
7.8378 | -87.6586
7.8382 | -87.6593
7.8385 | -87.66
7.8388 | -87.6601
7.8391 | -87.6604
7.8395 | -87.6615
7.8398 | -87.6622
7.8401 | -87.6622
7.8404 | -87.6627
7.8407 | -87.6633
7.8411 | -87.6641
7.8414 | -87.6647
7.8417 | -87.665
7.842 | -87.6657
7.8423 | -87.6659
7.8427 | -87.6658
7.843 | -87.666
7.8433 | -87.6665
7.8436 | -87.6667
7.844 | -87.6671
7.8443 | -87.6674
7.8446 | -87.668
7.8449 | -87.6678
7.8452 | -87.6684
7.8456 | -87.6692
7.8459 | -87.6705
7.8462 | -87.671
7.8465 | -87.6714
7.8468 | -87.6732
7.8472 | -87.675
7.8475 | -87.675
7.8478 | -87.675
7.8481 | -87.6752
7.8484 | -87.6751
7.8488 | -87.6748
7.8491 | -87.6745
7.8494 | -87.6742
7.8497 | -87.6742
7.8501 | -87.6748
7.8504 | -87.6752
7.8507 | -87.6758
7.851 | -87.6758
7.8513 | -87.6755
7.8517 | -87.6755
7.852 | -87.6757
7.8523 | -87.676
7.8526 | -87.6766
7.8529 | -87.6772
7.8533 | -87.6778
7.8536 | -87.6781
7.8539 | -87.678
7.8542 | -87.6789
7.8546 | -87.6795
7.8549 | -87.6792
7.8552 | -87.678
7.8555 | -87.6776
7.8558 | -87.678
7.8562 | -87.6781
7.8565 | -87.678
7.8568 | -87.6787
7.8571 | -87.6797
7.8574 | -87.6801
7.8578 | -87.6802
7.8581 | -87.6805
7.8584 | -87.6803
7.8587 | -87.6804
7.8591 | -87.6815
7.8594 | -87.6821
7.8597 | -87.6826
7.86 | -87.683
7.8603 | -87.6831
7.8607 | -87.6835
7.861 | -87.6836
7.8613 | -87.6835
7.8616 | -87.6832
7.8619 | -87.6835
7.8623 | -87.6836
7.8626 | -87.6845
7.8629 | -87.6848
7.8632 | -87.6854
7.8636 | -87.6868
7.8639 | -87.688
7.8642 | -87.6884
7.8645 | -87.6886
7.8648 | -87.6887
7.8652 | -87.6888
7.8655 | -87.689
7.8658 | -87.6895
7.8661 | -87.6897
7.8664 | -87.6902
7.8668 | -87.6907
7.8671 | -87.6904
7.8674 | -87.6908
7.8677 | -87.692
7.8681 | -87.6929
7.8684 | -87.6932
7.8687 | -87.6935
7.869 | -87.6938
7.8693 | -87.6942
7.8697 | -87.6945
7.87 | -87.6953
7.8703 | -87.6958
7.8706 | -87.696
7.8709 | -87.6964
7.8713 | -87.697
7.8716 | -87.6977
7.8719 | -87.6984
7.8722 | -87.6991
7.8725 | -87.6992
7.8729 | -87.7
7.8732 | -87.7004
7.8735 | -87.7006
7.8738 | -87.7011
7.8742 | -87.7011
7.8745 | -87.7011
7.8748 | -87.7017
7.8751 | -87.7021
7.8754 | -87.702
7.8758 | -87.7021
7.8761 | -87.7025
7.8764 | -87.7019
7.8767 | -87.7026
7.877 | -87.703
7.8774 | -87.7028
7.8777 | -87.7031
7.878 | -87.7046
7.8783 | -87.7063
7.8787 | -87.7066
7.879 | -87.707
7.8793 | -87.7069
7.8796 | -87.7068
7.8799 | -87.7072
7.8803 | -87.708
7.8806 | -87.7086
7.8809 | -87.7089
7.8812 | -87.7092
7.8815 | -87.7094
7.8819 | -87.7098
7.8822 | -87.7105
7.8825 | -87.7112
7.8828 | -87.7118
7.8832 | -87.7118
7.8835 | -87.7117
7.8838 | -87.712
7.8841 | -87.7121
7.8844 | -87.7126
7.8848 | -87.7128
7.8851 | -87.7128
7.8854 | -87.7133
7.8857 | -87.7144
7.886 | -87.7147
7.8864 | -87.7152
7.8867 | -87.7151
7.887 | -87.7149
7.8873 | -87.7152
7.8877 | -87.7155
7.888 | -87.7159
7.8883 | -87.7162
7.8886 | -87.7171
7.8889 | -87.717
7.8893 | -87.7168
7.8896 | -87.7175
7.8899 | -87.7178
7.8902 | -87.7184
7.8905 | -87.7192
7.8909 | -87.7183
7.8912 | -87.7166
7.8915 | -87.716
7.8918 | -87.7164
7.8921 | -87.7165
7.8925 | -87.7168
7.8928 | -87.7173
7.8931 | -87.7174
7.8934 | -87.7171
7.8938 | -87.7174
7.8941 | -87.7183
7.8944 | -87.7189
7.8947 | -87.7192
7.895 | -87.7194
7.8954 | -87.7192
7.8957 | -87.7194
7.896 | -87.7197
7.8963 | -87.7201
7.8966 | -87.7204
7.897 | -87.7207
7.8973 | -87.7213
7.8976 | -87.7212
7.8979 | -87.7212
7.8983 | -87.7213
7.8986 | -87.7217
7.8989 | -87.7222
7.8992 | -87.7222
7.8995 | -87.7225
7.8999 | -87.7229
7.9002 | -87.7234
7.9005 | -87.7252
7.9008 | -87.7265
7.9011 | -87.7266
7.9015 | -87.7266
7.9018 | -87.7264
7.9021 | -87.7261
7.9024 | -87.7261
7.9028 | -87.7269
7.9031 | -87.7274
7.9034 | -87.7277
7.9037 | -87.7278
7.904 | -87.7281
7.9044 | -87.7285
7.9047 | -87.7288
7.905 | -87.7292
7.9053 | -87.7294
7.9056 | -87.7298
7.906 | -87.7304
7.9063 | -87.7306
7.9066 | -87.7308
7.9069 | -87.731
7.9073 | -87.731
7.9076 | -87.7306
7.9079 | -87.7309
7.9082 | -87.7318
7.9085 | -87.732
7.9089 | -87.732
7.9092 | -87.7315
7.9095 | -87.7311
7.9098 | -87.7314
7.9101 | -87.7316
7.9105 | -87.7316
7.9108 | -87.7318
7.9111 | -87.732
7.9114 | -87.7329
7.9117 | -87.7344
7.9121 | -87.7354
7.9124 | -87.7359
7.9127 | -87.736
7.913 | -87.7362
7.9134 | -87.7365
7.9137 | -87.7369
7.914 | -87.7368
7.9143 | -87.7371
7.9146 | -87.7377
7.915 | -87.7381
7.9153 | -87.7383
7.9156 | -87.7391
7.9159 | -87.7394
7.9162 | -87.74
7.9166 | -87.7416
7.9169 | -87.7427
7.9172 | -87.7432
7.9175 | -87.7431
7.9179 | -87.7436
7.9182 | -87.7442
7.9185 | -87.7448
7.9188 | -87.7449
7.9191 | -87.7455
7.9195 | -87.7463
7.9198 | -87.7463
7.9201 | -87.7462
7.9204 | -87.746
7.9207 | -87.7464
7.9211 | -87.7464
7.9214 | -87.7463
7.9217 | -87.7468
7.922 | -87.7468
7.9224 | -87.7473
7.9227 | -87.7475
7.923 | -87.7475
7.9233 | -87.7481
7.9236 | -87.7491
7.924 | -87.7497
7.9243 | -87.7502
7.9246 | -87.7506
7.9249 | -87.751
7.9252 | -87.7514
7.9256 | -87.7518
7.9259 | -87.7521
7.9262 | -87.7534
7.9265 | -87.7535
7.9269 | -87.7552
7.9272 | -87.7563
7.9275 | -87.7567
7.9278 | -87.7582
7.9281 | -87.7591
7.9285 | -87.7596
7.9288 | -87.7601
7.9291 | -87.7606
7.9294 | -87.7616
7.9297 | -87.7618
7.9301 | -87.7616
7.9304 | -87.7612
7.9307 | -87.7611
7.931 | -87.7617
7.9314 | -87.7619
7.9317 | -87.7618
7.932 | -87.7618
7.9323 | -87.762
7.9326 | -87.7624
7.933 | -87.7624
7.9333 | -87.7623
7.9336 | -87.7627
7.9339 | -87.763
7.9342 | -87.7631
7.9346 | -87.7633
7.9349 | -87.7638
7.9352 | -87.7647
7.9355 | -87.7653
7.9359 | -87.765
7.9362 | -87.7653
7.9365 | -87.7656
7.9368 | -87.7655
7.9371 | -87.7659
7.9375 | -87.7665
7.9378 | -87.7669
7.9381 | -87.7671
7.9384 | -87.7674
7.9387 | -87.768
7.9391 | -87.7685
7.9394 | -87.7685
7.9397 | -87.7685
7.94 | -87.7686
7.9403 | -87.7689
7.9407 | -87.769
7.941 | -87.7693
7.9413 | -87.7705
7.9416 | -87.772
7.942 | -87.7728
7.9423 | -87.7732
7.9426 | -87.7727
7.9429 | -87.7724
7.9432 | -87.7726
7.9436 | -87.773
7.9439 | -87.7735
7.9442 | -87.774
7.9445 | -87.7746
7.9448 | -87.7761
7.9452 | -87.7762
7.9455 | -87.7777
7.9458 | -87.7782
7.9461 | -87.778
7.9465 | -87.7787
7.9468 | -87.7804
7.9471 | -87.7811
7.9474 | -87.7813
7.9477 | -87.7817
7.9481 | -87.7816
7.9484 | -87.782
7.9487 | -87.7828
7.949 | -87.7828
7.9493 | -87.7829
7.9497 | -87.7835
7.95 | -87.7839
7.9503 | -87.7842
7.9506 | -87.7848
7.951 | -87.7849
7.9513 | -87.7853
7.9516 | -87.7856
7.9519 | -87.786
7.9522 | -87.7862
7.9526 | -87.787
7.9529 | -87.7875
7.9532 | -87.7876
7.9535 | -87.7873
7.9538 | -87.7878
7.9542 | -87.7881
7.9545 | -87.7884
7.9548 | -87.7887
7.9551 | -87.789
7.9555 | -87.7899
7.9558 | -87.7908
7.9561 | -87.7912
7.9564 | -87.7917
7.9567 | -87.7918
7.9571 | -87.7918
7.9574 | -87.7924
7.9577 | -87.793
7.958 | -87.7933
7.9583 | -87.7942
7.9587 | -87.796
7.959 | -87.7967
7.9593 | -87.797
7.9596 | -87.7975
7.96 | -87.7974
7.9603 | -87.798
7.9606 | -87.7985
7.9609 | -87.7988
7.9612 | -87.7991
7.9616 | -87.799
7.9619 | -87.7991
7.9622 | -87.7995
7.9625 | -87.7995
7.9628 | -87.7997
7.9632 | -87.8008
7.9635 | -87.8025
7.9638 | -87.8035
7.9641 | -87.8048
7.9644 | -87.8053
7.9648 | -87.8051
7.9651 | -87.8051
7.9654 | -87.8055
7.9657 | -87.8061
7.9661 | -87.8066
7.9664 | -87.807
7.9667 | -87.8072
7.967 | -87.8073
7.9673 | -87.8074
7.9677 | -87.8076
7.968 | -87.8079
7.9683 | -87.8085
7.9686 | -87.8089
7.9689 | -87.8096
7.9693 | -87.8104
7.9696 | -87.8109
7.9699 | -87.8114
7.9702 | -87.812
7.9706 | -87.8121
7.9709 | -87.812
7.9712 | -87.8128
7.9715 | -87.8135
7.9718 | -87.8138
7.9722 | -87.8143
7.9725 | -87.8146
7.9728 | -87.8153
7.9731 | -87.816
7.9734 | -87.8164
7.9738 | -87.8174
7.9741 | -87.8188
7.9744 | -87.8201
7.9747 | -87.8206
7.9751 | -87.8212
7.9754 | -87.8216
7.9757 | -87.822
7.976 | -87.8221
7.9763 | -87.8222
7.9767 | -87.8226
7.977 | -87.8231
7.9773 | -87.8228
7.9776 | -87.8227
7.9779 | -87.8227
7.9783 | -87.823
7.9786 | -87.8236
7.9789 | -87.824
7.9792 | -87.8245
7.9796 | -87.8249
7.9799 | -87.8249
7.9802 | -87.8254
7.9805 | -87.8261
7.9808 | -87.827
7.9812 | -87.8276
7.9815 | -87.8278
7.9818 | -87.8276
7.9821 | -87.8279
7.9824 | -87.8288
7.9828 | -87.8293
7.9831 | -87.8298
7.9834 | -87.8306
7.9837 | -87.8309
7.984 | -87.8313
7.9844 | -87.8315
7.9847 | -87.832
7.985 | -87.8329
7.9853 | -87.8337
7.9857 | -87.8341
7.986 | -87.8349
7.9863 | -87.8351
7.9866 | -87.8354
7.9869 | -87.8364
7.9873 | -87.8371
7.9876 | -87.8377
7.9879 | -87.8382
7.9882 | -87.8388
7.9885 | -87.8391
7.9889 | -87.8405
7.9892 | -87.8415
7.9895 | -87.842
7.9898 | -87.8424
7.9902 | -87.8429
7.9905 | -87.8431
7.9908 | -87.8434
7.9911 | -87.8438
7.9914 | -87.845
7.9918 | -87.8458
7.9921 | -87.8469
7.9924 | -87.8475
7.9927 | -87.8463
7.993 | -87.8456
7.9934 | -87.8461
7.9937 | -87.8467
7.994 | -87.8474
7.9943 | -87.8485
7.9947 | -87.8498
7.995 | -87.8512
7.9953 | -87.8526
7.9956 | -87.8531
7.9959 | -87.8535
7.9963 | -87.8541
7.9966 | -87.8543
7.9969 | -87.8541
7.9972 | -87.8546
7.9975 | -87.8556
7.9979 | -87.8562
7.9982 | -87.8567
7.9985 | -87.8571
7.9988 | -87.8578
7.9992 | -87.8586
7.9995 | -87.8591
7.9998 | -87.859
8.0001 | -87.8583
8.0004 | -87.8578
8.0008 | -87.8586
8.0011 | -87.8592
8.0014 | -87.8598
8.0017 | -87.8602
8.002 | -87.8606
8.0024 | -87.8612
8.0027 | -87.8617
8.003 | -87.8625
8.0033 | -87.863
8.0037 | -87.8637
8.004 | -87.8638
8.0043 | -87.8637
8.0046 | -87.8642
8.0049 | -87.8645
8.0053 | -87.8648
8.0056 | -87.8653
8.0059 | -87.866
8.0062 | -87.8668
8.0065 | -87.8671
8.0069 | -87.8674
8.0072 | -87.8679
8.0075 | -87.8685
8.0078 | -87.8688
8.0081 | -87.869
8.0085 | -87.8696
8.0088 | -87.8701
8.0091 | -87.8712
8.0094 | -87.8718
8.0098 | -87.8722
8.0101 | -87.8732
8.0104 | -87.874
8.0107 | -87.8748
8.011 | -87.8759
8.0114 | -87.8768
8.0117 | -87.877
8.012 | -87.8776
8.0123 | -87.8787
8.0126 | -87.8793
8.013 | -87.8797
8.0133 | -87.8806
8.0136 | -87.8815
8.0139 | -87.8822
8.0143 | -87.8825
8.0146 | -87.8826
8.0149 | -87.8832
8.0152 | -87.8839
8.0155 | -87.8844
8.0159 | -87.8847
8.0162 | -87.8856
8.0165 | -87.8862
8.0168 | -87.8859
8.0171 | -87.8871
8.0175 | -87.8878
8.0178 | -87.8881
8.0181 | -87.8887
8.0184 | -87.8895
8.0188 | -87.8911
8.0191 | -87.892
8.0194 | -87.8927
8.0197 | -87.8934
8.02 | -87.8944
8.0204 | -87.8948
8.0207 | -87.8953
8.021 | -87.8963
8.0213 | -87.8968
8.0216 | -87.8972
8.022 | -87.897
8.0223 | -87.896
8.0226 | -87.8952
8.0229 | -87.8944
8.0233 | -87.8939
8.0236 | -87.894
8.0239 | -87.8948
8.0242 | -87.8954
8.0245 | -87.8958
8.0249 | -87.8967
8.0252 | -87.8975
8.0255 | -87.8976
8.0258 | -87.8984
8.0261 | -87.8992
8.0265 | -87.8997
8.0268 | -87.9002
8.0271 | -87.9017
8.0274 | -87.9024
8.0277 | -87.9029
8.0281 | -87.9035
8.0284 | -87.9039
8.0287 | -87.9051
8.029 | -87.9064
8.0294 | -87.9069
8.0297 | -87.9077
8.03 | -87.9085
8.0303 | -87.909
8.0306 | -87.9091
8.031 | -87.9103
8.0313 | -87.9116
8.0316 | -87.9125
8.0319 | -87.9135
8.0322 | -87.9142
8.0326 | -87.9149
8.0329 | -87.9157
8.0332 | -87.9164
8.0335 | -87.9171
8.0339 | -87.9179
8.0342 | -87.9187
8.0345 | -87.9189
8.0348 | -87.9197
8.0351 | -87.9202
8.0355 | -87.9208
8.0358 | -87.9216
8.0361 | -87.9231
8.0364 | -87.9244
8.0367 | -87.9246
8.0371 | -87.9243
8.0374 | -87.9246
8.0377 | -87.9254
8.038 | -87.9264
8.0384 | -87.9272
8.0387 | -87.928
8.039 | -87.9289
8.0393 | -87.9299
8.0396 | -87.9297
8.04 | -87.9299
8.0403 | -87.9302
8.0406 | -87.931
8.0409 | -87.9318
8.0412 | -87.9325
8.0416 | -87.9335
8.0419 | -87.9341
8.0422 | -87.9341
8.0425 | -87.9343
8.0429 | -87.935
8.0432 | -87.9353
8.0435 | -87.9352
8.0438 | -87.9351
8.0441 | -87.9356
8.0445 | -87.935
8.0448 | -87.935
8.0451 | -87.9363
8.0454 | -87.9372
8.0457 | -87.9382
8.0461 | -87.9387
8.0464 | -87.9399
8.0467 | -87.9413
8.047 | -87.9424
8.0474 | -87.943
8.0477 | -87.9436
8.048 | -87.9443
8.0483 | -87.9452
8.0486 | -87.9455
8.049 | -87.9461
8.0493 | -87.9467
8.0496 | -87.947
8.0499 | -87.9486
8.0502 | -87.95
8.0506 | -87.952
8.0509 | -87.9534
8.0512 | -87.9542
8.0515 | -87.9548
8.0518 | -87.9552
8.0522 | -87.956
8.0525 | -87.9576
8.0528 | -87.9583
8.0531 | -87.9589
8.0535 | -87.9595
8.0538 | -87.9601
8.0541 | -87.9607
8.0544 | -87.9612
8.0547 | -87.9617
8.0551 | -87.9623
8.0554 | -87.9637
8.0557 | -87.9638
8.056 | -87.963
8.0563 | -87.9633
8.0567 | -87.9638
8.057 | -87.965
8.0573 | -87.9661
8.0576 | -87.9668
8.058 | -87.9675
8.0583 | -87.968
8.0586 | -87.9684
8.0589 | -87.9673
8.0592 | -87.9673
8.0596 | -87.9676
8.0599 | -87.9682
8.0602 | -87.969
8.0605 | -87.9699
8.0608 | -87.9709
8.0612 | -87.9714
8.0615 | -87.9726
8.0618 | -87.9731
8.0621 | -87.9719
8.0625 | -87.9724
8.0628 | -87.9726
8.0631 | -87.9729
8.0634 | -87.9741
8.0637 | -87.975
8.0641 | -87.9755
8.0644 | -87.9759
8.0647 | -87.9763
8.065 | -87.977
8.0653 | -87.9781
8.0657 | -87.979
8.066 | -87.9799
8.0663 | -87.9808
8.0666 | -87.9817
8.067 | -87.9821
8.0673 | -87.9836
8.0676 | -87.9853
8.0679 | -87.9862
8.0682 | -87.987
8.0686 | -87.9885
8.0689 | -87.989
8.0692 | -87.9889
8.0695 | -87.9889
8.0698 | -87.9894
8.0702 | -87.9895
8.0705 | -87.99
8.0708 | -87.9908
8.0711 | -87.9914
8.0715 | -87.9922
8.0718 | -87.9933
8.0721 | -87.9941
8.0724 | -87.9945
8.0727 | -87.9951
8.0731 | -87.9956
8.0734 | -87.9962
8.0737 | -87.9976
8.074 | -87.9985
8.0743 | -87.9988
8.0747 | -87.9992
8.075 | -87.9998
8.0753 | -88.0008
8.0756 | -88.0015
8.0759 | -88.0019
8.0763 | -88.0021
8.0766 | -88.0025
8.0769 | -88.003
8.0772 | -88.0037
8.0776 | -88.0042
8.0779 | -88.0048
8.0782 | -88.0054
8.0785 | -88.0064
8.0788 | -88.0072
8.0792 | -88.0078
8.0795 | -88.0083
8.0798 | -88.009
8.0801 | -88.0096
8.0804 | -88.01
8.0808 | -88.0113
8.0811 | -88.012
8.0814 | -88.0127
8.0817 | -88.0132
8.0821 | -88.0129
8.0824 | -88.0125
8.0827 | -88.0133
8.083 | -88.0136
8.0833 | -88.0138
8.0837 | -88.0142
8.084 | -88.0137
8.0843 | -88.0141
8.0846 | -88.0141
8.0849 | -88.0145
8.0853 | -88.0144
8.0856 | -88.0154
8.0859 | -88.0165
8.0862 | -88.018
8.0866 | -88.0189
8.0869 | -88.0189
8.0872 | -88.0189
8.0875 | -88.0191
8.0878 | -88.0192
8.0882 | -88.02
8.0885 | -88.0206
8.0888 | -88.0214
8.0891 | -88.0219
8.0894 | -88.0224
8.0898 | -88.0234
8.0901 | -88.0257
8.0904 | -88.0265
8.0907 | -88.0271
8.0911 | -88.028
8.0914 | -88.0285
8.0917 | -88.0294
8.092 | -88.0303
8.0923 | -88.0309
8.0927 | -88.0313
8.093 | -88.0316
8.0933 | -88.032
8.0936 | -88.0327
8.0939 | -88.0333
8.0943 | -88.0337
8.0946 | -88.0338
8.0949 | -88.0345
8.0952 | -88.0363
8.0955 | -88.0372
8.0959 | -88.0378
8.0962 | -88.039
8.0965 | -88.04
8.0968 | -88.0406
8.0972 | -88.0412
8.0975 | -88.0418
8.0978 | -88.0426
8.0981 | -88.0428
8.0984 | -88.0429
8.0988 | -88.0431
8.0991 | -88.0432
8.0994 | -88.0432
8.0997 | -88.0435
8.1 | -88.0441
8.1004 | -88.0451
8.1007 | -88.0456
8.101 | -88.0464
8.1013 | -88.0469
8.1017 | -88.0476
8.102 | -88.0481
8.1023 | -88.0481
8.1026 | -88.0489
8.1029 | -88.0494
8.1033 | -88.0504
8.1036 | -88.0511
8.1039 | -88.052
8.1042 | -88.0529
8.1045 | -88.0535
8.1049 | -88.0542
8.1052 | -88.055
8.1055 | -88.0553
8.1058 | -88.0554
8.1062 | -88.0554
8.1065 | -88.057
8.1068 | -88.0582
8.1071 | -88.0589
8.1074 | -88.0603
8.1078 | -88.0604
8.1081 | -88.0612
8.1084 | -88.0621
8.1087 | -88.0627
8.109 | -88.0638
8.1094 | -88.064
8.1097 | -88.0647
8.11 | -88.0652
8.1103 | -88.0657
8.1107 | -88.0656
8.111 | -88.0658
8.1113 | -88.0663
8.1116 | -88.0667
8.1119 | -88.0665
8.1123 | -88.0662
8.1126 | -88.0661
8.1129 | -88.0657
8.1132 | -88.0661
8.1135 | -88.0673
8.1139 | -88.0685
8.1142 | -88.0696
8.1145 | -88.0705
8.1148 | -88.0714
8.1152 | -88.0718
8.1155 | -88.0722
8.1158 | -88.0726
8.1161 | -88.0732
8.1164 | -88.0732
8.1168 | -88.0739
8.1171 | -88.0749
8.1174 | -88.0757
8.1177 | -88.0764
8.118 | -88.0772
8.1184 | -88.0774
8.1187 | -88.0776
8.119 | -88.0778
8.1193 | -88.0791
8.1197 | -88.0795
8.12 | -88.0799
8.1203 | -88.0807
8.1206 | -88.0809
8.1209 | -88.0814
8.1213 | -88.0818
8.1216 | -88.0824
8.1219 | -88.083
8.1222 | -88.0835
8.1225 | -88.084
8.1229 | -88.0851
8.1232 | -88.0858
8.1235 | -88.086
8.1238 | -88.0863
8.1241 | -88.0871
8.1245 | -88.0878
8.1248 | -88.0884
8.1251 | -88.0889
8.1254 | -88.0891
8.1258 | -88.0889
8.1261 | -88.0893
8.1264 | -88.0891
8.1267 | -88.0888
8.127 | -88.0888
8.1274 | -88.0892
8.1277 | -88.089
8.128 | -88.0892
8.1283 | -88.0901
8.1286 | -88.0904
8.129 | -88.0904
8.1293 | -88.0906
8.1296 | -88.0912
8.1299 | -88.0913
8.1303 | -88.0912
8.1306 | -88.0909
8.1309 | -88.0909
8.1312 | -88.0897
8.1315 | -88.0895
8.1319 | -88.0895
8.1322 | -88.089
8.1325 | -88.0884
8.1328 | -88.0871
8.1331 | -88.0855
8.1335 | -88.0846
8.1338 | -88.0846
8.1341 | -88.0846
8.1344 | -88.0837
8.1348 | -88.0836
8.1351 | -88.0834
8.1354 | -88.0826
8.1357 | -88.0805
8.136 | -88.08
8.1364 | -88.0799
8.1367 | -88.0799
8.137 | -88.0796
8.1373 | -88.0798
8.1376 | -88.0801
8.138 | -88.0795
8.1383 | -88.0793
8.1386 | -88.0794
8.1389 | -88.0796
8.1393 | -88.0794
8.1396 | -88.0794
8.1399 | -88.0786
8.1402 | -88.0789
8.1405 | -88.0789
8.1409 | -88.0784
8.1412 | -88.0775
8.1415 | -88.0771
8.1418 | -88.0776
8.1421 | -88.0777
8.1425 | -88.0775
8.1428 | -88.0766
8.1431 | -88.0759
8.1434 | -88.0762
8.1438 | -88.076
8.1441 | -88.076
8.1444 | -88.0762
8.1447 | -88.0765
8.145 | -88.0769
8.1454 | -88.0772
8.1457 | -88.0775
8.146 | -88.0774
8.1463 | -88.0769
8.1466 | -88.076
8.147 | -88.0758
8.1473 | -88.0759
8.1476 | -88.0765
8.1479 | -88.0757
8.1482 | -88.0751
8.1486 | -88.0746
8.1489 | -88.0743
8.1492 | -88.0741
8.1495 | -88.0745
8.1499 | -88.0752
8.1502 | -88.0754
8.1505 | -88.0752
8.1508 | -88.0744
8.1511 | -88.074
8.1515 | -88.074
8.1518 | -88.0742
8.1521 | -88.0741
8.1524 | -88.0739
8.1527 | -88.0739
8.1531 | -88.074
8.1534 | -88.0739
8.1537 | -88.0735
8.154 | -88.0737
8.1544 | -88.0743
8.1547 | -88.0746
8.155 | -88.0749
8.1553 | -88.0749
8.1556 | -88.0747
8.156 | -88.0745
8.1563 | -88.0746
8.1566 | -88.0747
8.1569 | -88.075
8.1572 | -88.0746
8.1576 | -88.0742
8.1579 | -88.0744
8.1582 | -88.0745
8.1585 | -88.0745
8.1589 | -88.0746
8.1592 | -88.0744
8.1595 | -88.0743
8.1598 | -88.0747
8.1601 | -88.0748
8.1605 | -88.0748
8.1608 | -88.0755
8.1611 | -88.076
8.1614 | -88.076
8.1617 | -88.0759
8.1621 | -88.0755
8.1624 | -88.0753
8.1627 | -88.0757
8.163 | -88.076
8.1634 | -88.0757
8.1637 | -88.076
8.164 | -88.0755
8.1643 | -88.0752
8.1646 | -88.0753
8.165 | -88.0751
8.1653 | -88.075
8.1656 | -88.0752
8.1659 | -88.0752
8.1662 | -88.0751
8.1666 | -88.0752
8.1669 | -88.0761
8.1672 | -88.0769
8.1675 | -88.0772
8.1678 | -88.0771
8.1682 | -88.0768
8.1685 | -88.0765
8.1688 | -88.0762
8.1691 | -88.0757
8.1695 | -88.0755
8.1698 | -88.0755
8.1701 | -88.0755
8.1704 | -88.0749
8.1707 | -88.0745
8.1711 | -88.0743
8.1714 | -88.0744
8.1717 | -88.0741
8.172 | -88.0743
8.1723 | -88.0755
8.1727 | -88.0763
8.173 | -88.0761
8.1733 | -88.0765
8.1736 | -88.0769
8.174 | -88.0769
8.1743 | -88.0767
8.1746 | -88.0769
8.1749 | -88.0773
8.1752 | -88.0758
8.1756 | -88.0754
8.1759 | -88.0753
8.1762 | -88.0749
8.1765 | -88.0746
8.1768 | -88.0738
8.1772 | -88.0744
8.1775 | -88.0747
8.1778 | -88.0743
8.1781 | -88.0742
8.1785 | -88.0744
8.1788 | -88.0745
8.1791 | -88.0745
8.1794 | -88.0741
8.1797 | -88.0738
8.1801 | -88.0738
8.1804 | -88.0734
8.1807 | -88.0726
8.181 | -88.0722
8.1813 | -88.0719
8.1817 | -88.072
8.182 | -88.0718
8.1823 | -88.0712
8.1826 | -88.0712
8.183 | -88.0714
8.1833 | -88.0718
8.1836 | -88.0723
8.1839 | -88.0724
8.1842 | -88.0717
8.1846 | -88.0709
8.1849 | -88.0708
8.1852 | -88.0705
8.1855 | -88.0707
8.1858 | -88.0709
8.1862 | -88.0679
8.1865 | -88.0665
8.1868 | -88.0658
8.1871 | -88.0655
8.1875 | -88.0649
8.1878 | -88.0646
8.1881 | -88.0643
8.1884 | -88.0638
8.1887 | -88.0636
8.1891 | -88.0629
8.1894 | -88.0621
8.1897 | -88.0621
8.19 | -88.0614
8.1903 | -88.061
8.1907 | -88.0607
8.191 | -88.0603
8.1913 | -88.0599
8.1916 | -88.0607
8.1919 | -88.0602
8.1923 | -88.0599
8.1926 | -88.0596
8.1929 | -88.0593
8.1932 | -88.059
8.1936 | -88.0583
8.1939 | -88.0579
8.1942 | -88.0575
8.1945 | -88.0575
8.1948 | -88.057
8.1952 | -88.0564
8.1955 | -88.0558
8.1958 | -88.0552
8.1961 | -88.0549
8.1964 | -88.0548
8.1968 | -88.0545
8.1971 | -88.0543
8.1974 | -88.0543
8.1977 | -88.0541
8.1981 | -88.0538
8.1984 | -88.0536
8.1987 | -88.0532
8.199 | -88.0529
8.1993 | -88.0524
8.1997 | -88.052
8.2 | -88.0511
8.2003 | -88.0501
8.2006 | -88.0501
8.2009 | -88.0507
8.2013 | -88.0502
8.2016 | -88.0498
8.2019 | -88.0492
8.2022 | -88.0487
8.2026 | -88.0479
8.2029 | -88.0476
8.2032 | -88.0468
8.2035 | -88.0458
8.2038 | -88.0446
8.2042 | -88.044
8.2045 | -88.0434
8.2048 | -88.0427
8.2051 | -88.0421
8.2054 | -88.0417
8.2058 | -88.0413
8.2061 | -88.0409
8.2064 | -88.0408
8.2067 | -88.0403
8.2071 | -88.0395
8.2074 | -88.0391
8.2077 | -88.0385
8.208 | -88.0377
8.2083 | -88.0373
8.2087 | -88.037
8.209 | -88.0366
8.2093 | -88.0362
8.2096 | -88.0359
8.2099 | -88.0352
8.2103 | -88.034
8.2106 | -88.0338
8.2109 | -88.0339
8.2112 | -88.0339
8.2116 | -88.0334
8.2119 | -88.0327
8.2122 | -88.032
8.2125 | -88.0323
8.2128 | -88.0329
8.2132 | -88.0332
8.2135 | -88.0321
8.2138 | -88.0315
8.2141 | -88.0309
8.2144 | -88.0303
8.2148 | -88.0301
8.2151 | -88.0296
8.2154 | -88.0289
8.2157 | -88.0284
8.216 | -88.0278
8.2164 | -88.0273
8.2167 | -88.0268
8.217 | -88.0265
8.2173 | -88.0262
8.2177 | -88.0259
8.218 | -88.0252
8.2183 | -88.0248
8.2186 | -88.0245
8.2189 | -88.0243
8.2193 | -88.0248
8.2196 | -88.0259
8.2199 | -88.0261
8.2202 | -88.0252
8.2205 | -88.0236
8.2209 | -88.0228
8.2212 | -88.0222
8.2215 | -88.0214
8.2218 | -88.0207
8.2222 | -88.02
8.2225 | -88.0194
8.2228 | -88.0187
8.2231 | -88.018
8.2234 | -88.0174
8.2238 | -88.0168
8.2241 | -88.0163
8.2244 | -88.0158
8.2247 | -88.0163
8.225 | -88.0163
8.2254 | -88.0155
8.2257 | -88.0147
8.226 | -88.0143
8.2263 | -88.0142
8.2267 | -88.0136
8.227 | -88.0129
8.2273 | -88.0123
8.2276 | -88.0117
8.2279 | -88.0113
8.2283 | -88.0111
8.2286 | -88.0107
8.2289 | -88.0104
8.2292 | -88.0098
8.2295 | -88.009
8.2299 | -88.0084
8.2302 | -88.0077
8.2305 | -88.0071
8.2308 | -88.0065
8.2312 | -88.0059
8.2315 | -88.0052
8.2318 | -88.0046
8.2321 | -88.0043
8.2324 | -88.0025
8.2328 | -88.0011
8.2331 | -88.0006
8.2334 | -88.0001
8.2337 | -87.9996
8.234 | -87.9989
8.2344 | -87.998
8.2347 | -87.9976
8.235 | -87.997
8.2353 | -87.9959
8.2356 | -87.9949
8.236 | -87.9944
8.2363 | -87.9938
8.2366 | -87.9933
8.2369 | -87.9925
8.2373 | -87.9911
8.2376 | -87.9897
8.2379 | -87.9886
8.2382 | -87.9871
8.2385 | -87.9854
8.2389 | -87.9832
8.2392 | -87.9819
8.2395 | -87.9812
8.2398 | -87.9806
8.2401 | -87.9796
8.2405 | -87.9779
8.2408 | -87.9768
8.2411 | -87.9761
8.2414 | -87.9757
8.2418 | -87.9735
8.2421 | -87.9729
8.2424 | -87.9723
8.2427 | -87.9721
8.243 | -87.972
8.2434 | -87.9714
8.2437 | -87.9708
8.244 | -87.9705
8.2443 | -87.97
8.2446 | -87.9692
8.245 | -87.9686
8.2453 | -87.9686
8.2456 | -87.9696
8.2459 | -87.9696
8.2463 | -87.9694
8.2466 | -87.9684
8.2469 | -87.9673
8.2472 | -87.9663
8.2475 | -87.9659
8.2479 | -87.9653
8.2482 | -87.9645
8.2485 | -87.9641
8.2488 | -87.9635
8.2491 | -87.9633
8.2495 | -87.963
8.2498 | -87.9625
8.2501 | -87.9618
8.2504 | -87.9612
8.2508 | -87.9601
8.2511 | -87.9592
8.2514 | -87.9584
8.2517 | -87.9574
8.252 | -87.9565
8.2524 | -87.9568
8.2527 | -87.9558
8.253 | -87.9545
8.2533 | -87.9535
8.2536 | -87.9526
8.254 | -87.9518
8.2543 | -87.9503
8.2546 | -87.9487
8.2549 | -87.9479
8.2553 | -87.947
8.2556 | -87.9461
8.2559 | -87.9452
8.2562 | -87.9444
8.2565 | -87.9438
8.2569 | -87.9428
8.2572 | -87.9422
8.2575 | -87.9422
8.2578 | -87.9424
8.2581 | -87.9418
8.2585 | -87.941
8.2588 | -87.9405
8.2591 | -87.9397
8.2594 | -87.9388
8.2597 | -87.9378
8.2601 | -87.9368
8.2604 | -87.936
8.2607 | -87.9341
8.261 | -87.9329
8.2614 | -87.9322
8.2617 | -87.9318
8.262 | -87.9311
8.2623 | -87.9305
8.2626 | -87.9298
8.263 | -87.9288
8.2633 | -87.9278
8.2636 | -87.9266
8.2639 | -87.9249
8.2642 | -87.9246
8.2646 | -87.9242
8.2649 | -87.9237
8.2652 | -87.923
8.2655 | -87.9225
8.2659 | -87.9219
8.2662 | -87.921
8.2665 | -87.9202
8.2668 | -87.9192
8.2671 | -87.918
8.2675 | -87.9171
8.2678 | -87.9163
8.2681 | -87.9156
8.2684 | -87.9149
8.2687 | -87.9124
8.2691 | -87.9107
8.2694 | -87.9089
8.2697 | -87.9071
8.27 | -87.9062
8.2704 | -87.9052
8.2707 | -87.9042
8.271 | -87.9031
8.2713 | -87.9023
8.2716 | -87.9015
8.272 | -87.9012
8.2723 | -87.9007
8.2726 | -87.9001
8.2729 | -87.899
8.2732 | -87.8982
8.2736 | -87.897
8.2739 | -87.8951
8.2742 | -87.8943
8.2745 | -87.8933
8.2749 | -87.8923
8.2752 | -87.8913
8.2755 | -87.8903
8.2758 | -87.8895
8.2761 | -87.8891
8.2765 | -87.8886
8.2768 | -87.8879
8.2771 | -87.8873
8.2774 | -87.8865
8.2777 | -87.8856
8.2781 | -87.8842
8.2784 | -87.8831
8.2787 | -87.8827
8.279 | -87.8818
8.2794 | -87.8804
8.2797 | -87.8795
8.28 | -87.8788
8.2803 | -87.8782
8.2806 | -87.8774
8.281 | -87.8766
8.2813 | -87.8756
8.2816 | -87.8742
8.2819 | -87.873
8.2822 | -87.8721
8.2826 | -87.8712
8.2829 | -87.8705
8.2832 | -87.8701
8.2835 | -87.8692
8.2838 | -87.8683
8.2842 | -87.8674
8.2845 | -87.8667
8.2848 | -87.8661
8.2851 | -87.8651
8.2855 | -87.8644
8.2858 | -87.8643
8.2861 | -87.8637
8.2864 | -87.8625
8.2867 | -87.8615
8.2871 | -87.8607
8.2874 | -87.8599
8.2877 | -87.8606
8.288 | -87.8605
8.2883 | -87.86
8.2887 | -87.8591
8.289 | -87.8584
8.2893 | -87.8577
8.2896 | -87.8563
8.29 | -87.8551
8.2903 | -87.8547
8.2906 | -87.8543
8.2909 | -87.8538
8.2912 | -87.8529
8.2916 | -87.8521
8.2919 | -87.8518
8.2922 | -87.8504
8.2925 | -87.848
8.2928 | -87.8458
8.2932 | -87.8443
8.2935 | -87.8433
8.2938 | -87.8423
8.2941 | -87.8411
8.2945 | -87.8402
8.2948 | -87.8397
8.2951 | -87.8392
8.2954 | -87.8383
8.2957 | -87.8374
8.2961 | -87.8368
8.2964 | -87.836
8.2967 | -87.8351
8.297 | -87.8342
8.2973 | -87.8338
8.2977 | -87.8327
8.298 | -87.8315
8.2983 | -87.8303
8.2986 | -87.8294
8.299 | -87.8288
8.2993 | -87.8279
8.2996 | -87.8272
8.2999 | -87.8264
8.3002 | -87.8255
8.3006 | -87.825
8.3009 | -87.8239
8.3012 | -87.8225
8.3015 | -87.8214
8.3018 | -87.8206
8.3022 | -87.8199
8.3025 | -87.8193
8.3028 | -87.8184
8.3031 | -87.8176
8.3034 | -87.8168
8.3038 | -87.816
8.3041 | -87.8153
8.3044 | -87.8144
8.3047 | -87.8136
8.3051 | -87.8129
8.3054 | -87.8115
8.3057 | -87.8096
8.306 | -87.8085
8.3063 | -87.8074
8.3067 | -87.8062
8.307 | -87.8056
8.3073 | -87.8048
8.3076 | -87.8032
8.3079 | -87.8018
8.3083 | -87.8007
8.3086 | -87.8
8.3089 | -87.7989
8.3092 | -87.7983
8.3096 | -87.7974
8.3099 | -87.7967
8.3102 | -87.7966
8.3105 | -87.7957
8.3108 | -87.7949
8.3112 | -87.7943
8.3115 | -87.7931
8.3118 | -87.7918
8.3121 | -87.7904
8.3124 | -87.789
8.3128 | -87.7882
8.3131 | -87.7881
8.3134 | -87.7875
8.3137 | -87.7868
8.3141 | -87.7865
8.3144 | -87.7858
8.3147 | -87.785
8.315 | -87.784
8.3153 | -87.7828
8.3157 | -87.7819
8.316 | -87.7814
8.3163 | -87.78
8.3166 | -87.7791
8.3169 | -87.7779
8.3173 | -87.7764
8.3176 | -87.7763
8.3179 | -87.7771
8.3182 | -87.7772
8.3186 | -87.7758
8.3189 | -87.7748
8.3192 | -87.7739
8.3195 | -87.7728
8.3198 | -87.7713
8.3202 | -87.7701
8.3205 | -87.7688
8.3208 | -87.7679
8.3211 | -87.7672
8.3214 | -87.766
8.3218 | -87.7639
8.3221 | -87.7625
8.3224 | -87.7614
8.3227 | -87.7603
8.3231 | -87.7596
8.3234 | -87.7588
8.3237 | -87.7586
8.324 | -87.7576
8.3243 | -87.7556
8.3247 | -87.7543
8.325 | -87.7536
8.3253 | -87.7535
8.3256 | -87.7526
8.3259 | -87.7512
8.3263 | -87.7504
8.3266 | -87.7496
8.3269 | -87.7482
8.3272 | -87.7464
8.3275 | -87.7446
8.3279 | -87.7433
8.3282 | -87.7422
8.3285 | -87.7414
8.3288 | -87.7411
8.3292 | -87.7406
8.3295 | -87.7397
8.3298 | -87.7388
8.3301 | -87.7381
8.3304 | -87.7367
8.3308 | -87.7352
8.3311 | -87.7342
8.3314 | -87.7335
8.3317 | -87.7327
8.332 | -87.7319
8.3324 | -87.7307
8.3327 | -87.7295
8.333 | -87.7286
8.3333 | -87.728
8.3337 | -87.7266
8.334 | -87.7253
8.3343 | -87.7246
8.3346 | -87.7244
8.3349 | -87.7235
8.3353 | -87.7229
8.3356 | -87.7225
8.3359 | -87.7216
8.3362 | -87.7212
8.3365 | -87.7205
8.3369 | -87.7196
8.3372 | -87.7189
8.3375 | -87.7182
8.3378 | -87.7174
8.3382 | -87.7168
8.3385 | -87.7162
8.3388 | -87.7157
8.3391 | -87.7148
8.3394 | -87.7141
8.3398 | -87.7133
8.3401 | -87.7122
8.3404 | -87.7109
8.3407 | -87.7102
8.341 | -87.71
8.3414 | -87.7094
8.3417 | -87.7086
8.342 | -87.7077
8.3423 | -87.707
8.3427 | -87.7066
8.343 | -87.7062
8.3433 | -87.7053
8.3436 | -87.7048
8.3439 | -87.7045
8.3443 | -87.7038
8.3446 | -87.703
8.3449 | -87.7027
8.3452 | -87.7023
8.3455 | -87.7017
8.3459 | -87.7011
8.3462 | -87.7004
8.3465 | -87.6995
8.3468 | -87.699
8.3471 | -87.6991
8.3475 | -87.6989
8.3478 | -87.6983
8.3481 | -87.6975
8.3484 | -87.6969
8.3488 | -87.6962
8.3491 | -87.6957
8.3494 | -87.6954
8.3497 | -87.6946
8.35 | -87.6936
8.3504 | -87.6927
8.3507 | -87.6914
8.351 | -87.691
8.3513 | -87.6906
8.3516 | -87.6898
8.352 | -87.689
8.3523 | -87.6882
8.3526 | -87.6874
8.3529 | -87.6868
8.3533 | -87.6869
8.3536 | -87.6869
8.3539 | -87.6874
8.3542 | -87.6877
8.3545 | -87.6877
8.3549 | -87.6877
8.3552 | -87.6869
8.3555 | -87.6862
8.3558 | -87.6857
8.3561 | -87.6859
8.3565 | -87.6854
8.3568 | -87.6839
8.3571 | -87.6824
8.3574 | -87.6814
8.3578 | -87.6796
8.3581 | -87.6782
8.3584 | -87.6769
8.3587 | -87.6759
8.359 | -87.675
8.3594 | -87.674
8.3597 | -87.674
8.36 | -87.6744
8.3603 | -87.6742
8.3606 | -87.6728
8.361 | -87.6712
8.3613 | -87.6702
8.3616 | -87.6694
8.3619 | -87.6689
8.3623 | -87.669
8.3626 | -87.6686
8.3629 | -87.6679
8.3632 | -87.6666
8.3635 | -87.666
8.3639 | -87.6651
8.3642 | -87.6648
8.3645 | -87.6651
8.3648 | -87.665
8.3651 | -87.6643
8.3655 | -87.6639
8.3658 | -87.6631
8.3661 | -87.6624
8.3664 | -87.6617
8.3668 | -87.6611
8.3671 | -87.6603
8.3674 | -87.6598
8.3677 | -87.6596
8.368 | -87.6605
8.3684 | -87.6603
8.3687 | -87.6593
8.369 | -87.6588
8.3693 | -87.6585
8.3696 | -87.6574
8.37 | -87.6562
8.3703 | -87.6555
8.3706 | -87.6552
8.3709 | -87.655
8.3712 | -87.655
8.3716 | -87.6548
8.3719 | -87.6542
8.3722 | -87.6539
8.3725 | -87.6536
8.3729 | -87.6528
8.3732 | -87.6522
8.3735 | -87.6516
8.3738 | -87.6516
8.3741 | -87.6512
8.3745 | -87.6505
8.3748 | -87.65
8.3751 | -87.6493
8.3754 | -87.6492
8.3757 | -87.6486
8.3761 | -87.6481
8.3764 | -87.647
8.3767 | -87.6475
8.377 | -87.648
8.3774 | -87.6465
8.3777 | -87.6461
8.378 | -87.6456
8.3783 | -87.645
8.3786 | -87.6446
8.379 | -87.6445
8.3793 | -87.6445
8.3796 | -87.6443
8.3799 | -87.6433
8.3802 | -87.6428
8.3806 | -87.6422
8.3809 | -87.6412
8.3812 | -87.6406
8.3815 | -87.6403
8.3819 | -87.64
8.3822 | -87.6398
8.3825 | -87.6389
8.3828 | -87.6386
8.3831 | -87.6379
8.3835 | -87.637
8.3838 | -87.6368
8.3841 | -87.6359
8.3844 | -87.6354
8.3847 | -87.6349
8.3851 | -87.6347
8.3854 | -87.6342
8.3857 | -87.6333
8.386 | -87.6325
8.3864 | -87.632
8.3867 | -87.6316
8.387 | -87.6307
8.3873 | -87.6293
8.3876 | -87.6287
8.388 | -87.6287
8.3883 | -87.6283
8.3886 | -87.6276
8.3889 | -87.627
8.3892 | -87.6261
8.3896 | -87.6251
8.3899 | -87.6244
8.3902 | -87.6238
8.3905 | -87.6234
8.3909 | -87.6231
8.3912 | -87.6228
8.3915 | -87.6221
8.3918 | -87.6213
8.3921 | -87.6208
8.3925 | -87.62
8.3928 | -87.6195
8.3931 | -87.6192
8.3934 | -87.6187
8.3937 | -87.6181
8.3941 | -87.6176
8.3944 | -87.6169
8.3947 | -87.6166
8.395 | -87.6166
8.3954 | -87.616
8.3957 | -87.6157
8.396 | -87.6156
8.3963 | -87.6153
8.3966 | -87.615
8.397 | -87.6146
8.3973 | -87.6143
8.3976 | -87.614
8.3979 | -87.6135
8.3982 | -87.6132
8.3986 | -87.6128
8.3989 | -87.6122
8.3992 | -87.6113
8.3995 | -87.6111
8.3998 | -87.6105
8.4002 | -87.6099
8.4005 | -87.6102
8.4008 | -87.6106
8.4011 | -87.6102
8.4015 | -87.6099
8.4018 | -87.6097
8.4021 | -87.6093
8.4024 | -87.609
8.4027 | -87.6087
8.4031 | -87.6083
8.4034 | -87.6082
8.4037 | -87.6081
8.404 | -87.6078
8.4043 | -87.6074
8.4047 | -87.6069
8.405 | -87.6065
8.4053 | -87.606
8.4056 | -87.6055
8.406 | -87.6051
8.4063 | -87.605
8.4066 | -87.6048
8.4069 | -87.6046
8.4072 | -87.6044
8.4076 | -87.6043
8.4079 | -87.604
8.4082 | -87.6036
8.4085 | -87.6034
8.4088 | -87.6026
8.4092 | -87.6025
8.4095 | -87.6022
8.4098 | -87.602
8.4101 | -87.6012
8.4105 | -87.6006
8.4108 | -87.6005
8.4111 | -87.6001
8.4114 | -87.5998
8.4117 | -87.5995
8.4121 | -87.5993
8.4124 | -87.599
8.4127 | -87.5987
8.413 | -87.5982
8.4133 | -87.5983
8.4137 | -87.5991
8.414 | -87.599
8.4143 | -87.5987
8.4146 | -87.5985
8.415 | -87.5982
8.4153 | -87.5981
8.4156 | -87.598
8.4159 | -87.5977
8.4162 | -87.5974
8.4166 | -87.5972
8.4169 | -87.5968
8.4172 | -87.5963
8.4175 | -87.5958
8.4178 | -87.5951
8.4182 | -87.5943
8.4185 | -87.5938
8.4188 | -87.5934
8.4191 | -87.5932
8.4194 | -87.5931
8.4198 | -87.5932
8.4201 | -87.5941
8.4204 | -87.5943
8.4207 | -87.5939
8.4211 | -87.5935
8.4214 | -87.5934
8.4217 | -87.5932
8.422 | -87.5928
8.4223 | -87.5921
8.4227 | -87.5919
8.423 | -87.5917
8.4233 | -87.5917
8.4236 | -87.5916
8.4239 | -87.5913
8.4243 | -87.5911
8.4246 | -87.591
8.4249 | -87.5904
8.4252 | -87.59
8.4256 | -87.5898
8.4259 | -87.5897
8.4262 | -87.5895
8.4265 | -87.5899
8.4268 | -87.59
8.4272 | -87.5903
8.4275 | -87.5901
8.4278 | -87.59
8.4281 | -87.5896
8.4284 | -87.5893
8.4288 | -87.5886
8.4291 | -87.5883
8.4294 | -87.5881
8.4297 | -87.5877
8.4301 | -87.5871
8.4304 | -87.5871
8.4307 | -87.587
8.431 | -87.5867
8.4313 | -87.5864
8.4317 | -87.5864
8.432 | -87.5863
8.4323 | -87.586
8.4326 | -87.5864
8.4329 | -87.5866
8.4333 | -87.5866
8.4336 | -87.5866
8.4339 | -87.5865
8.4342 | -87.5864
8.4346 | -87.5863
8.4349 | -87.5862
8.4352 | -87.5861
8.4355 | -87.586
8.4358 | -87.5861
8.4362 | -87.586
8.4365 | -87.5867
8.4368 | -87.586
8.4371 | -87.5857
8.4374 | -87.5855
8.4378 | -87.585
8.4381 | -87.5852
8.4384 | -87.5851
8.4387 | -87.5851
8.439 | -87.5847
8.4394 | -87.5843
8.4397 | -87.5841
8.44 | -87.584
8.4403 | -87.584
8.4407 | -87.5838
8.441 | -87.5835
8.4413 | -87.5833
8.4416 | -87.5836
8.4419 | -87.5837
8.4423 | -87.584
8.4426 | -87.5844
8.4429 | -87.5834
8.4432 | -87.5833
8.4435 | -87.5837
8.4439 | -87.5838
8.4442 | -87.5833
8.4445 | -87.5828
8.4448 | -87.5828
8.4452 | -87.5829
8.4455 | -87.5831
8.4458 | -87.5835
8.4461 | -87.5837
8.4464 | -87.5837
8.4468 | -87.5833
8.4471 | -87.5827
8.4474 | -87.5825
8.4477 | -87.5829
8.448 | -87.5828
8.4484 | -87.5829
8.4487 | -87.5831
8.449 | -87.583
8.4493 | -87.5829
8.4497 | -87.5827
8.45 | -87.5829
8.4503 | -87.5832
8.4506 | -87.5834
8.4509 | -87.5836
8.4513 | -87.5837
8.4516 | -87.5836
8.4519 | -87.584
8.4522 | -87.5842
8.4525 | -87.5842
8.4529 | -87.5842
8.4532 | -87.5845
8.4535 | -87.5849
8.4538 | -87.5851
8.4542 | -87.585
8.4545 | -87.5846
8.4548 | -87.5846
8.4551 | -87.5852
8.4554 | -87.5852
8.4558 | -87.5853
8.4561 | -87.5859
8.4564 | -87.5861
8.4567 | -87.5859
8.457 | -87.5862
8.4574 | -87.5861
8.4577 | -87.5861
8.458 | -87.5864
8.4583 | -87.5873
8.4587 | -87.5888
8.459 | -87.5896
8.4593 | -87.5896
8.4596 | -87.5895
8.4599 | -87.5896
8.4603 | -87.5895
8.4606 | -87.5896
8.4609 | -87.5902
8.4612 | -87.5903
8.4615 | -87.5899
8.4619 | -87.5891
8.4622 | -87.5892
8.4625 | -87.5892
8.4628 | -87.5898
8.4632 | -87.5899
8.4635 | -87.59
8.4638 | -87.59
8.4641 | -87.59
8.4644 | -87.5904
8.4648 | -87.5903
8.4651 | -87.5906
8.4654 | -87.5912
8.4657 | -87.5906
8.466 | -87.5889
8.4664 | -87.5885
8.4667 | -87.5885
8.467 | -87.5887
8.4673 | -87.5886
8.4676 | -87.5889
8.468 | -87.589
8.4683 | -87.5889
8.4686 | -87.589
8.4689 | -87.5887
8.4693 | -87.5889
8.4696 | -87.5892
8.4699 | -87.5895
8.4702 | -87.5901
8.4705 | -87.5904
8.4709 | -87.5905
8.4712 | -87.5896
8.4715 | -87.589
8.4718 | -87.589
8.4721 | -87.5892
8.4725 | -87.5891
8.4728 | -87.589
8.4731 | -87.5885
8.4734 | -87.5878
8.4738 | -87.5881
8.4741 | -87.5884
8.4744 | -87.5888
8.4747 | -87.5889
8.475 | -87.5887
8.4754 | -87.5885
8.4757 | -87.5888
8.476 | -87.5895
8.4763 | -87.5898
8.4766 | -87.59
8.477 | -87.5898
8.4773 | -87.5896
8.4776 | -87.5897
8.4779 | -87.5906
8.4783 | -87.5915
8.4786 | -87.5919
8.4789 | -87.5919
8.4792 | -87.5921
8.4795 | -87.5924
8.4799 | -87.5927
8.4802 | -87.5929
8.4805 | -87.5931
8.4808 | -87.5932
8.4811 | -87.5935
8.4815 | -87.5941
8.4818 | -87.5944
8.4821 | -87.594
8.4824 | -87.5941
8.4828 | -87.5945
8.4831 | -87.5949
8.4834 | -87.5954
8.4837 | -87.5955
8.484 | -87.5957
8.4844 | -87.5957
8.4847 | -87.5959
8.485 | -87.5966
8.4853 | -87.5967
8.4856 | -87.5968
8.486 | -87.5971
8.4863 | -87.5974
8.4866 | -87.5977
8.4869 | -87.5978
8.4872 | -87.598
8.4876 | -87.5983
8.4879 | -87.5985
8.4882 | -87.599
8.4885 | -87.5993
8.4889 | -87.5995
8.4892 | -87.6004
8.4895 | -87.6016
8.4898 | -87.6024
8.4901 | -87.6026
8.4905 | -87.6029
8.4908 | -87.6033
8.4911 | -87.6035
8.4914 | -87.6038
8.4917 | -87.6042
8.4921 | -87.6051
8.4924 | -87.6054
8.4927 | -87.606
8.493 | -87.6065
8.4934 | -87.6067
8.4937 | -87.6068
8.494 | -87.6071
8.4943 | -87.6074
8.4946 | -87.6079
8.495 | -87.6083
8.4953 | -87.6084
8.4956 | -87.6084
8.4959 | -87.6084
8.4962 | -87.6088
8.4966 | -87.6092
8.4969 | -87.6089
8.4972 | -87.6091
8.4975 | -87.6094
8.4979 | -87.6097
8.4982 | -87.6098
8.4985 | -87.6098
8.4988 | -87.6101
8.4991 | -87.6106
8.4995 | -87.6115
8.4998 | -87.6123
8.5001 | -87.613
8.5004 | -87.6139
8.5007 | -87.6146
8.5011 | -87.6149
8.5014 | -87.6155
8.5017 | -87.6158
8.502 | -87.6162
8.5024 | -87.6167
8.5027 | -87.6169
8.503 | -87.6169
8.5033 | -87.6174
8.5036 | -87.6176
8.504 | -87.6179
8.5043 | -87.6181
8.5046 | -87.6186
8.5049 | -87.6187
8.5052 | -87.6186
8.5056 | -87.6182
8.5059 | -87.6181
8.5062 | -87.619
8.5065 | -87.6197
8.5069 | -87.6201
8.5072 | -87.6206
8.5075 | -87.6213
8.5078 | -87.6217
8.5081 | -87.6219
8.5085 | -87.6221
8.5088 | -87.6224
8.5091 | -87.6225
8.5094 | -87.6228
8.5097 | -87.6231
8.5101 | -87.6238
8.5104 | -87.6247
8.5107 | -87.6252
8.511 | -87.6255
8.5113 | -87.6256
8.5117 | -87.6262
8.512 | -87.6268
8.5123 | -87.6273
8.5126 | -87.6276
8.513 | -87.6279
8.5133 | -87.6284
8.5136 | -87.6295
8.5139 | -87.6303
8.5142 | -87.6307
8.5146 | -87.631
8.5149 | -87.6315
8.5152 | -87.6314
8.5155 | -87.6315
8.5158 | -87.6319
8.5162 | -87.6325
8.5165 | -87.6335
8.5168 | -87.6341
8.5171 | -87.6345
8.5175 | -87.6352
8.5178 | -87.636
8.5181 | -87.6363
8.5184 | -87.6367
8.5187 | -87.6371
8.5191 | -87.6375
8.5194 | -87.6378
8.5197 | -87.6381
8.52 | -87.639
8.5203 | -87.6394
8.5207 | -87.6402
8.521 | -87.6408
8.5213 | -87.6412
8.5216 | -87.6417
8.522 | -87.6423
8.5223 | -87.6423
8.5226 | -87.6424
8.5229 | -87.6422
8.5232 | -87.642
8.5236 | -87.6417
8.5239 | -87.6421
8.5242 | -87.6425
8.5245 | -87.6427
8.5248 | -87.6432
8.5252 | -87.6446
8.5255 | -87.6452
8.5258 | -87.6456
8.5261 | -87.6462
8.5265 | -87.6463
8.5268 | -87.646
8.5271 | -87.6461
8.5274 | -87.6466
8.5277 | -87.6471
8.5281 | -87.6479
8.5284 | -87.6485
8.5287 | -87.649
8.529 | -87.6493
8.5293 | -87.6496
8.5297 | -87.6498
8.53 | -87.6504
8.5303 | -87.6512
8.5306 | -87.6517
8.531 | -87.652
8.5313 | -87.6526
8.5316 | -87.6535
8.5319 | -87.6536
8.5322 | -87.654
8.5326 | -87.6546
8.5329 | -87.6549
8.5332 | -87.6552
8.5335 | -87.6555
8.5338 | -87.6563
8.5342 | -87.657
8.5345 | -87.6573
8.5348 | -87.6575
8.5351 | -87.6576
8.5354 | -87.6581
8.5358 | -87.6587
8.5361 | -87.659
8.5364 | -87.6593
8.5367 | -87.6598
8.5371 | -87.6604
8.5374 | -87.6611
8.5377 | -87.6617
8.538 | -87.6621
8.5383 | -87.6626
8.5387 | -87.6634
8.539 | -87.6637
8.5393 | -87.664
8.5396 | -87.6644
8.5399 | -87.665
8.5403 | -87.6649
8.5406 | -87.6652
8.5409 | -87.6656
8.5412 | -87.666
8.5416 | -87.6666
8.5419 | -87.6674
8.5422 | -87.668
8.5425 | -87.6685
8.5428 | -87.6691
8.5432 | -87.6695
8.5435 | -87.6703
8.5438 | -87.6714
8.5441 | -87.672
8.5444 | -87.6724
8.5448 | -87.6729
8.5451 | -87.6734
8.5454 | -87.674
8.5457 | -87.6745
8.5461 | -87.6746
8.5464 | -87.6753
8.5467 | -87.6761
8.547 | -87.6764
8.5473 | -87.6765
8.5477 | -87.677
8.548 | -87.6773
8.5483 | -87.6779
8.5486 | -87.6785
8.5489 | -87.6789
8.5493 | -87.6795
8.5496 | -87.6802
8.5499 | -87.6813
8.5502 | -87.6817
8.5506 | -87.6822
8.5509 | -87.6832
8.5512 | -87.684
8.5515 | -87.6848
8.5518 | -87.6854
8.5522 | -87.686
8.5525 | -87.6864
8.5528 | -87.6865
8.5531 | -87.6869
8.5534 | -87.6878
8.5538 | -87.6882
8.5541 | -87.6883
8.5544 | -87.6888
8.5547 | -87.6896
8.555 | -87.6904
8.5554 | -87.6916
8.5557 | -87.6925
8.556 | -87.6933
8.5563 | -87.6939
8.5567 | -87.6945
8.557 | -87.6953
8.5573 | -87.6955
8.5576 | -87.6953
8.5579 | -87.6958
8.5583 | -87.6965
8.5586 | -87.6972
8.5589 | -87.6971
8.5592 | -87.6975
8.5595 | -87.6981
8.5599 | -87.6986
8.5602 | -87.6991
8.5605 | -87.6999
8.5608 | -87.7006
8.5612 | -87.7006
8.5615 | -87.7008
8.5618 | -87.7013
8.5621 | -87.7014
8.5624 | -87.7021
8.5628 | -87.7028
8.5631 | -87.7035
8.5634 | -87.7046
8.5637 | -87.7056
8.564 | -87.7064
8.5644 | -87.707
8.5647 | -87.7077
8.565 | -87.7084
8.5653 | -87.709
8.5657 | -87.7095
8.566 | -87.71
8.5663 | -87.7106
8.5666 | -87.7108
8.5669 | -87.7114
8.5673 | -87.7122
8.5676 | -87.7143
8.5679 | -87.7161
8.5682 | -87.7166
8.5685 | -87.7171
8.5689 | -87.7176
8.5692 | -87.7182
8.5695 | -87.7189
8.5698 | -87.7195
8.5702 | -87.7202
8.5705 | -87.7211
8.5708 | -87.7218
8.5711 | -87.7225
8.5714 | -87.7232
8.5718 | -87.7244
8.5721 | -87.7252
8.5724 | -87.7257
8.5727 | -87.7263
8.573 | -87.7267
8.5734 | -87.7272
8.5737 | -87.7277
8.574 | -87.7283
8.5743 | -87.7289
8.5747 | -87.7301
8.575 | -87.7317
8.5753 | -87.7328
8.5756 | -87.733
8.5759 | -87.7339
8.5763 | -87.7344
8.5766 | -87.7349
8.5769 | -87.7353
8.5772 | -87.7357
8.5775 | -87.7365
8.5779 | -87.7372
8.5782 | -87.7378
8.5785 | -87.7382
8.5788 | -87.7384
8.5792 | -87.7392
8.5795 | -87.74
8.5798 | -87.7412
8.5801 | -87.7419
8.5804 | -87.7428
8.5808 | -87.7435
8.5811 | -87.7442
8.5814 | -87.7449
8.5817 | -87.7453
8.582 | -87.7461
8.5824 | -87.7467
8.5827 | -87.7474
8.583 | -87.7479
8.5833 | -87.7487
8.5836 | -87.7496
8.584 | -87.7503
8.5843 | -87.7514
8.5846 | -87.7522
8.5849 | -87.7525
8.5853 | -87.7534
8.5856 | -87.7543
8.5859 | -87.755
8.5862 | -87.7557
8.5865 | -87.7564
8.5869 | -87.7573
8.5872 | -87.758
8.5875 | -87.7586
8.5878 | -87.7597
8.5881 | -87.7606
8.5885 | -87.7613
8.5888 | -87.7622
8.5891 | -87.763
8.5894 | -87.7638
8.5898 | -87.7645
8.5901 | -87.7652
8.5904 | -87.7661
8.5907 | -87.7671
8.591 | -87.7679
8.5914 | -87.7685
8.5917 | -87.7692
8.592 | -87.77
8.5923 | -87.7703
8.5926 | -87.7706
8.593 | -87.7719
8.5933 | -87.7728
8.5936 | -87.7735
8.5939 | -87.7743
8.5943 | -87.7744
8.5946 | -87.7741
8.5949 | -87.7745
8.5952 | -87.7752
8.5955 | -87.7759
8.5959 | -87.7766
8.5962 | -87.7772
8.5965 | -87.778
8.5968 | -87.7787
8.5971 | -87.78
8.5975 | -87.7812
8.5978 | -87.7816
8.5981 | -87.782
8.5984 | -87.7825
8.5988 | -87.7838
8.5991 | -87.7858
8.5994 | -87.7879
8.5997 | -87.7896
8.6 | -87.7911
8.6004 | -87.7918
8.6007 | -87.7921
8.601 | -87.7928
8.6013 | -87.7937
8.6016 | -87.7942
8.602 | -87.7952
8.6023 | -87.7963
8.6026 | -87.7971
8.6029 | -87.7976
8.6032 | -87.7982
8.6036 | -87.7994
8.6039 | -87.8003
8.6042 | -87.8011
8.6045 | -87.802
8.6049 | -87.8032
8.6052 | -87.8042
8.6055 | -87.8044
8.6058 | -87.8049
8.6061 | -87.8056
8.6065 | -87.8059
8.6068 | -87.8062
8.6071 | -87.8072
8.6074 | -87.8077
8.6077 | -87.8083
8.6081 | -87.8093
8.6084 | -87.8102
8.6087 | -87.8109
8.609 | -87.8115
8.6094 | -87.8121
8.6097 | -87.812
8.61 | -87.8122
8.6103 | -87.813
8.6106 | -87.8138
8.611 | -87.8145
8.6113 | -87.8152
8.6116 | -87.8162
8.6119 | -87.8173
8.6122 | -87.8177
8.6126 | -87.819
8.6129 | -87.8197
8.6132 | -87.8204
8.6135 | -87.821
8.6139 | -87.8216
8.6142 | -87.8224
8.6145 | -87.8235
8.6148 | -87.8249
8.6151 | -87.8255
8.6155 | -87.826
8.6158 | -87.8268
8.6161 | -87.8276
8.6164 | -87.8284
8.6167 | -87.8293
8.6171 | -87.8301
8.6174 | -87.8309
8.6177 | -87.8325
8.618 | -87.8337
8.6184 | -87.8349
8.6187 | -87.8357
8.619 | -87.8368
8.6193 | -87.8376
8.6196 | -87.8383
8.62 | -87.8388
8.6203 | -87.8399
8.6206 | -87.8409
8.6209 | -87.8417
8.6212 | -87.8427
8.6216 | -87.844
8.6219 | -87.8452
8.6222 | -87.8463
8.6225 | -87.8474
8.6228 | -87.8493
8.6232 | -87.8508
8.6235 | -87.8513
8.6238 | -87.852
8.6241 | -87.8529
8.6245 | -87.8537
8.6248 | -87.8545
8.6251 | -87.8549
8.6254 | -87.8556
8.6257 | -87.8565
8.6261 | -87.8572
8.6264 | -87.858
8.6267 | -87.8588
8.627 | -87.8594
8.6273 | -87.8604
8.6277 | -87.8612
8.628 | -87.8621
8.6283 | -87.8633
8.6286 | -87.8642
8.629 | -87.8651
8.6293 | -87.8658
8.6296 | -87.8664
8.6299 | -87.8673
8.6302 | -87.8683
8.6306 | -87.8692
8.6309 | -87.87
8.6312 | -87.8706
8.6315 | -87.8712
8.6318 | -87.8719
8.6322 | -87.8727
8.6325 | -87.8735
8.6328 | -87.8743
8.6331 | -87.8753
8.6335 | -87.876
8.6338 | -87.8761
8.6341 | -87.8768
8.6344 | -87.8777
8.6347 | -87.8794
8.6351 | -87.8818
8.6354 | -87.8834
8.6357 | -87.8842
8.636 | -87.885
8.6363 | -87.886
8.6367 | -87.8865
8.637 | -87.8868
8.6373 | -87.8873
8.6376 | -87.8884
8.638 | -87.8896
8.6383 | -87.8919
8.6386 | -87.8934
8.6389 | -87.894
8.6392 | -87.8942
8.6396 | -87.895
8.6399 | -87.8962
8.6402 | -87.897
8.6405 | -87.8983
8.6408 | -87.899
8.6412 | -87.8998
8.6415 | -87.9011
8.6418 | -87.9036
8.6421 | -87.9051
8.6425 | -87.9066
8.6428 | -87.9077
8.6431 | -87.9084
8.6434 | -87.9091
8.6437 | -87.9102
8.6441 | -87.9114
8.6444 | -87.9124
8.6447 | -87.9134
8.645 | -87.9144
8.6453 | -87.9153
8.6457 | -87.9156
8.646 | -87.9161
8.6463 | -87.9172
8.6466 | -87.9181
8.647 | -87.9193
8.6473 | -87.9202
8.6476 | -87.921
8.6479 | -87.9215
8.6482 | -87.9223
8.6486 | -87.9232
8.6489 | -87.924
8.6492 | -87.9254
8.6495 | -87.9269
8.6498 | -87.9283
8.6502 | -87.9295
8.6505 | -87.9306
8.6508 | -87.9314
8.6511 | -87.932
8.6514 | -87.9328
8.6518 | -87.9333
8.6521 | -87.9341
8.6524 | -87.935
8.6527 | -87.9361
8.6531 | -87.937
8.6534 | -87.9379
8.6537 | -87.9388
8.654 | -87.9397
8.6543 | -87.9406
8.6547 | -87.9411
8.655 | -87.9415
8.6553 | -87.9416
8.6556 | -87.9424
8.6559 | -87.9433
8.6563 | -87.9443
8.6566 | -87.9455
8.6569 | -87.9467
8.6572 | -87.9476
8.6576 | -87.9487
8.6579 | -87.9496
8.6582 | -87.9506
8.6585 | -87.9517
8.6588 | -87.9525
8.6592 | -87.9539
8.6595 | -87.9549
8.6598 | -87.9562
8.6601 | -87.9571
8.6604 | -87.9583
8.6608 | -87.9587
8.6611 | -87.9593
8.6614 | -87.9602
8.6617 | -87.9605
8.6621 | -87.9612
8.6624 | -87.9619
8.6627 | -87.9625
8.663 | -87.9638
8.6633 | -87.9657
8.6637 | -87.9668
8.664 | -87.968
8.6643 | -87.969
8.6646 | -87.9701
8.6649 | -87.9721
8.6653 | -87.9733
8.6656 | -87.9741
8.6659 | -87.9756
8.6662 | -87.9768
8.6666 | -87.9784
8.6669 | -87.9794
8.6672 | -87.9798
8.6675 | -87.9802
8.6678 | -87.9811
8.6682 | -87.9824
8.6685 | -87.9836
8.6688 | -87.9852
8.6691 | -87.9865
8.6694 | -87.9876
8.6698 | -87.9883
8.6701 | -87.9891
8.6704 | -87.9903
8.6707 | -87.9916
8.671 | -87.9924
8.6714 | -87.9932
8.6717 | -87.994
8.672 | -87.995
8.6723 | -87.9959
8.6727 | -87.9965
8.673 | -87.9974
8.6733 | -87.9982
8.6736 | -87.9991
8.6739 | -88
8.6743 | -88.0011
8.6746 | -88.0021
8.6749 | -88.0029
8.6752 | -88.0039
8.6755 | -88.0051
8.6759 | -88.0064
8.6762 | -88.0075
8.6765 | -88.0089
8.6768 | -88.0096
8.6772 | -88.0106
8.6775 | -88.0114
8.6778 | -88.0124
8.6781 | -88.0134
8.6784 | -88.0142
8.6788 | -88.0146
8.6791 | -88.0158
8.6794 | -88.0169
8.6797 | -88.0179
8.68 | -88.0189
8.6804 | -88.0201
8.6807 | -88.0219
8.681 | -88.0233
8.6813 | -88.0244
8.6817 | -88.0257
8.682 | -88.0271
8.6823 | -88.028
8.6826 | -88.0291
8.6829 | -88.0301
8.6833 | -88.0312
8.6836 | -88.0323
8.6839 | -88.033
8.6842 | -88.0329
8.6845 | -88.0338
8.6849 | -88.0355
8.6852 | -88.0367
8.6855 | -88.0378
8.6858 | -88.039
8.6862 | -88.04
8.6865 | -88.0411
8.6868 | -88.0423
8.6871 | -88.0434
8.6874 | -88.0445
8.6878 | -88.0454
8.6881 | -88.0464
8.6884 | -88.0474
8.6887 | -88.0482
8.689 | -88.0499
8.6894 | -88.0513
8.6897 | -88.0526
8.69 | -88.0537
8.6903 | -88.0548
8.6907 | -88.0562
8.691 | -88.0573
8.6913 | -88.0584
8.6916 | -88.0592
8.6919 | -88.0601
8.6923 | -88.0611
8.6926 | -88.062
8.6929 | -88.0626
8.6932 | -88.0629
8.6935 | -88.0642
8.6939 | -88.0657
8.6942 | -88.0666
8.6945 | -88.0677
8.6948 | -88.069
8.6951 | -88.0708
8.6955 | -88.0727
8.6958 | -88.0744
8.6961 | -88.0758
8.6964 | -88.0773
8.6968 | -88.0778
8.6971 | -88.0786
8.6974 | -88.0796
8.6977 | -88.0804
8.698 | -88.0813
8.6984 | -88.0825
8.6987 | -88.0839
8.699 | -88.0852
8.6993 | -88.0865
8.6996 | -88.0874
8.7 | -88.0888
8.7003 | -88.0899
8.7006 | -88.0909
8.7009 | -88.0921
8.7013 | -88.093
8.7016 | -88.0934
8.7019 | -88.0943
8.7022 | -88.0958
8.7025 | -88.0971
8.7029 | -88.0982
8.7032 | -88.0991
8.7035 | -88.1005
8.7038 | -88.1019
8.7041 | -88.1029
8.7045 | -88.1039
8.7048 | -88.1048
8.7051 | -88.106
8.7054 | -88.1074
8.7058 | -88.1086
8.7061 | -88.1094
8.7064 | -88.1105
8.7067 | -88.1116
8.707 | -88.1124
8.7074 | -88.1138
8.7077 | -88.1152
8.708 | -88.1167
8.7083 | -88.1177
8.7086 | -88.1193
8.709 | -88.1206
8.7093 | -88.1216
8.7096 | -88.1229
8.7099 | -88.1245
8.7103 | -88.1257
8.7106 | -88.1274
8.7109 | -88.129
8.7112 | -88.1301
8.7115 | -88.1312
8.7119 | -88.1329
8.7122 | -88.1344
8.7125 | -88.1352
8.7128 | -88.1366
8.7131 | -88.1386
8.7135 | -88.1401
8.7138 | -88.1408
8.7141 | -88.1413
8.7144 | -88.1424
8.7148 | -88.1439
8.7151 | -88.1456
8.7154 | -88.1469
8.7157 | -88.1483
8.716 | -88.1508
8.7164 | -88.153
8.7167 | -88.1549
8.717 | -88.1569
8.7173 | -88.159
8.7176 | -88.1617
8.718 | -88.1647
8.7183 | -88.1669
8.7186 | -88.1688
8.7189 | -88.1702
8.7192 | -88.1723
8.7196 | -88.1737
8.7199 | -88.1751
8.7202 | -88.1775
8.7205 | -88.1806
8.7209 | -88.1823
8.7212 | -88.1834
8.7215 | -88.1844
8.7218 | -88.1857
8.7221 | -88.1872
8.7225 | -88.1888
8.7228 | -88.1908
8.7231 | -88.192
8.7234 | -88.1926
8.7237 | -88.1935
8.7241 | -88.1945
8.7244 | -88.1956
8.7247 | -88.1967
8.725 | -88.1982
8.7254 | -88.1993
8.7257 | -88.2001
8.726 | -88.2012
8.7263 | -88.2028
8.7266 | -88.2048
8.727 | -88.2071
8.7273 | -88.2082
8.7276 | -88.2093
8.7279 | -88.2104
8.7282 | -88.2123
8.7286 | -88.2139
8.7289 | -88.2144
8.7292 | -88.2153
8.7295 | -88.2164
8.7299 | -88.2172
8.7302 | -88.2182
8.7305 | -88.2193
8.7308 | -88.2203
8.7311 | -88.2211
8.7315 | -88.2219
8.7318 | -88.2237
8.7321 | -88.2258
8.7324 | -88.227
8.7327 | -88.2281
8.7331 | -88.2291
8.7334 | -88.23
8.7337 | -88.2323
8.734 | -88.2338
8.7344 | -88.235
8.7347 | -88.2364
8.735 | -88.2373
8.7353 | -88.2386
8.7356 | -88.2395
8.736 | -88.2406
8.7363 | -88.2416
8.7366 | -88.2424
8.7369 | -88.2442
8.7372 | -88.2455
8.7376 | -88.2461
8.7379 | -88.2463
8.7382 | -88.247
8.7385 | -88.248
8.7388 | -88.249
8.7392 | -88.2498
8.7395 | -88.251
8.7398 | -88.2524
8.7401 | -88.2534
8.7405 | -88.2544
8.7408 | -88.2546
8.7411 | -88.2548
8.7414 | -88.2555
8.7417 | -88.2567
8.7421 | -88.2577
8.7424 | -88.259
8.7427 | -88.26
8.743 | -88.2609
8.7433 | -88.2619
8.7437 | -88.2634
8.744 | -88.2647
8.7443 | -88.2657
8.7446 | -88.267
8.745 | -88.2679
8.7453 | -88.2682
8.7456 | -88.2687
8.7459 | -88.2698
8.7462 | -88.271
8.7466 | -88.272
8.7469 | -88.2732
8.7472 | -88.2741
8.7475 | -88.2751
8.7478 | -88.2762
8.7482 | -88.2772
8.7485 | -88.2783
8.7488 | -88.2798
8.7491 | -88.2812
8.7495 | -88.2822
8.7498 | -88.2834
8.7501 | -88.2849
8.7504 | -88.286
8.7507 | -88.2875
8.7511 | -88.2889
8.7514 | -88.2902
8.7517 | -88.2916
8.752 | -88.2928
8.7523 | -88.2937
8.7527 | -88.2944
8.753 | -88.2954
8.7533 | -88.2963
8.7536 | -88.2972
8.754 | -88.2983
8.7543 | -88.2995
8.7546 | -88.3005
8.7549 | -88.3015
8.7552 | -88.3025
8.7556 | -88.3042
8.7559 | -88.3062
8.7562 | -88.3077
8.7565 | -88.309
8.7568 | -88.3094
8.7572 | -88.3105
8.7575 | -88.3119
8.7578 | -88.3131
8.7581 | -88.3145
8.7585 | -88.3152
8.7588 | -88.3162
8.7591 | -88.3174
8.7594 | -88.3186
8.7597 | -88.3197
8.7601 | -88.3208
8.7604 | -88.3222
8.7607 | -88.3231
8.761 | -88.3241
8.7613 | -88.3248
8.7617 | -88.3257
8.762 | -88.3267
8.7623 | -88.3275
8.7626 | -88.3287
8.763 | -88.3312
8.7633 | -88.3326
8.7636 | -88.3337
8.7639 | -88.3349
8.7642 | -88.3358
8.7646 | -88.3372
8.7649 | -88.3383
8.7652 | -88.3392
8.7655 | -88.3395
8.7658 | -88.3404
8.7662 | -88.3415
8.7665 | -88.3426
8.7668 | -88.3436
8.7671 | -88.3449
8.7674 | -88.3461
8.7678 | -88.3476
8.7681 | -88.3484
8.7684 | -88.3495
8.7687 | -88.351
8.7691 | -88.3521
8.7694 | -88.3531
8.7697 | -88.3541
8.77 | -88.3547
8.7703 | -88.3558
8.7707 | -88.3573
8.771 | -88.3583
8.7713 | -88.3594
8.7716 | -88.3605
8.7719 | -88.3612
8.7723 | -88.3619
8.7726 | -88.3636
8.7729 | -88.3653
8.7732 | -88.3665
8.7736 | -88.3676
8.7739 | -88.3684
8.7742 | -88.3696
8.7745 | -88.3716
8.7748 | -88.3739
8.7752 | -88.376
8.7755 | -88.3774
8.7758 | -88.3785
8.7761 | -88.3798
8.7764 | -88.3807
8.7768 | -88.3818
8.7771 | -88.3831
8.7774 | -88.3842
8.7777 | -88.3855
8.7781 | -88.3869
8.7784 | -88.3879
8.7787 | -88.3895
8.779 | -88.3905
8.7793 | -88.3907
8.7797 | -88.3917
8.78 | -88.3928
8.7803 | -88.394
8.7806 | -88.3954
8.7809 | -88.3968
8.7813 | -88.3981
8.7816 | -88.3991
8.7819 | -88.4003
8.7822 | -88.4015
8.7826 | -88.4029
8.7829 | -88.404
8.7832 | -88.405
8.7835 | -88.4061
8.7838 | -88.4075
8.7842 | -88.4089
8.7845 | -88.41
8.7848 | -88.4112
8.7851 | -88.4124
8.7854 | -88.4133
8.7858 | -88.4146
8.7861 | -88.4158
8.7864 | -88.4168
8.7867 | -88.4178
8.787 | -88.4187
8.7874 | -88.4199
8.7877 | -88.4208
8.788 | -88.4218
8.7883 | -88.4229
8.7887 | -88.4241
8.789 | -88.4254
8.7893 | -88.4264
8.7896 | -88.4279
8.7899 | -88.4296
8.7903 | -88.4308
8.7906 | -88.4319
8.7909 | -88.4331
8.7912 | -88.4337
8.7915 | -88.4345
8.7919 | -88.4355
8.7922 | -88.437
8.7925 | -88.4379
8.7928 | -88.4395
8.7932 | -88.4411
8.7935 | -88.4429
8.7938 | -88.4446
8.7941 | -88.446
8.7944 | -88.4471
8.7948 | -88.448
8.7951 | -88.4489
8.7954 | -88.4499
8.7957 | -88.4509
8.796 | -88.452
8.7964 | -88.4533
8.7967 | -88.4546
8.797 | -88.456
8.7973 | -88.4573
8.7977 | -88.4586
8.798 | -88.4595
8.7983 | -88.4604
8.7986 | -88.4611
8.7989 | -88.4618
8.7993 | -88.4627
8.7996 | -88.4636
8.7999 | -88.4643
8.8002 | -88.4658
8.8005 | -88.4672
8.8009 | -88.468
8.8012 | -88.4688
8.8015 | -88.4698
8.8018 | -88.4707
8.8022 | -88.4716
8.8025 | -88.4728
8.8028 | -88.4738
8.8031 | -88.4748
8.8034 | -88.4761
8.8038 | -88.4774
8.8041 | -88.4782
8.8044 | -88.4791
8.8047 | -88.4796
8.805 | -88.4804
8.8054 | -88.4814
8.8057 | -88.4824
8.806 | -88.4836
8.8063 | -88.4847
8.8066 | -88.4857
8.807 | -88.4867
8.8073 | -88.4876
8.8076 | -88.4887
8.8079 | -88.4901
8.8083 | -88.4912
8.8086 | -88.492
8.8089 | -88.4929
8.8092 | -88.4939
8.8095 | -88.4947
8.8099 | -88.4964
8.8102 | -88.4977
8.8105 | -88.4987
8.8108 | -88.4997
8.8111 | -88.5011
8.8115 | -88.5031
8.8118 | -88.5041
8.8121 | -88.505
8.8124 | -88.5063
8.8128 | -88.5074
8.8131 | -88.5086
8.8134 | -88.5097
8.8137 | -88.5106
8.814 | -88.5114
8.8144 | -88.5121
8.8147 | -88.5131
8.815 | -88.5141
8.8153 | -88.515
8.8156 | -88.5154
8.816 | -88.5159
8.8163 | -88.517
8.8166 | -88.5182
8.8169 | -88.5196
8.8173 | -88.521
8.8176 | -88.522
8.8179 | -88.523
8.8182 | -88.5241
8.8185 | -88.5252
8.8189 | -88.5262
8.8192 | -88.5273
8.8195 | -88.528
8.8198 | -88.5289
8.8201 | -88.5299
8.8205 | -88.5309
8.8208 | -88.532
8.8211 | -88.5333
8.8214 | -88.5342
8.8218 | -88.5351
8.8221 | -88.536
8.8224 | -88.5371
8.8227 | -88.5382
8.823 | -88.5393
8.8234 | -88.5402
8.8237 | -88.541
8.824 | -88.5422
8.8243 | -88.5444
8.8246 | -88.5459
8.825 | -88.5469
8.8253 | -88.5478
8.8256 | -88.5485
8.8259 | -88.5492
8.8263 | -88.5503
8.8266 | -88.5513
8.8269 | -88.552
8.8272 | -88.5529
8.8275 | -88.5542
8.8279 | -88.5554
8.8282 | -88.5565
8.8285 | -88.558
8.8288 | -88.5591
8.8291 | -88.5602
8.8295 | -88.5619
8.8298 | -88.564
8.8301 | -88.5657
8.8304 | -88.5677
8.8308 | -88.5701
8.8311 | -88.5725
8.8314 | -88.5743
8.8317 | -88.5757
8.832 | -88.5767
8.8324 | -88.5784
8.8327 | -88.5805
8.833 | -88.5829
8.8333 | -88.5847
8.8336 | -88.5859
8.834 | -88.5876
8.8343 | -88.5893
8.8346 | -88.5902
8.8349 | -88.5911
8.8352 | -88.5923
8.8356 | -88.5932
8.8359 | -88.5942
8.8362 | -88.5951
8.8365 | -88.596
8.8369 | -88.5965
8.8372 | -88.5973
8.8375 | -88.5985
8.8378 | -88.5996
8.8381 | -88.6007
8.8385 | -88.6017
8.8388 | -88.6026
8.8391 | -88.6036
8.8394 | -88.6051
8.8397 | -88.6064
8.8401 | -88.6077
8.8404 | -88.6087
8.8407 | -88.6097
8.841 | -88.6106
8.8414 | -88.6117
8.8417 | -88.6128
8.842 | -88.6139
8.8423 | -88.6149
8.8426 | -88.616
8.843 | -88.6169
8.8433 | -88.6177
8.8436 | -88.619
8.8439 | -88.6202
8.8442 | -88.6212
8.8446 | -88.6223
8.8449 | -88.6232
8.8452 | -88.624
8.8455 | -88.6252
8.8459 | -88.6262
8.8462 | -88.6276
8.8465 | -88.6288
8.8468 | -88.63
8.8471 | -88.631
8.8475 | -88.6318
8.8478 | -88.6334
8.8481 | -88.6352
8.8484 | -88.6361
8.8487 | -88.6371
8.8491 | -88.6381
8.8494 | -88.6389
8.8497 | -88.6398
8.85 | -88.6408
8.8504 | -88.6418
8.8507 | -88.6429
8.851 | -88.6436
8.8513 | -88.6444
8.8516 | -88.6455
8.852 | -88.6465
8.8523 | -88.6474
8.8526 | -88.6482
8.8529 | -88.6494
8.8532 | -88.6505
8.8536 | -88.6517
8.8539 | -88.6526
8.8542 | -88.6537
8.8545 | -88.6545
8.8548 | -88.6564
8.8552 | -88.6576
8.8555 | -88.6587
8.8558 | -88.6597
8.8561 | -88.6605
8.8565 | -88.6617
8.8568 | -88.6629
8.8571 | -88.6637
8.8574 | -88.665
8.8577 | -88.6667
8.8581 | -88.6685
8.8584 | -88.6693
8.8587 | -88.6698
8.859 | -88.6704
8.8593 | -88.6713
8.8597 | -88.6722
8.86 | -88.6727
8.8603 | -88.6735
8.8606 | -88.6743
8.861 | -88.6754
8.8613 | -88.6767
8.8616 | -88.6777
8.8619 | -88.6791
8.8622 | -88.6805
8.8626 | -88.6815
8.8629 | -88.6826
8.8632 | -88.6857
8.8635 | -88.6876
8.8638 | -88.6899
8.8642 | -88.6922
8.8645 | -88.6931
8.8648 | -88.6943
8.8651 | -88.6954
8.8655 | -88.6965
8.8658 | -88.6977
8.8661 | -88.6985
8.8664 | -88.6992
8.8667 | -88.6997
8.8671 | -88.7004
8.8674 | -88.701
8.8677 | -88.7015
8.868 | -88.7023
8.8683 | -88.7035
8.8687 | -88.7051
8.869 | -88.7054
8.8693 | -88.7061
8.8696 | -88.7072
8.87 | -88.7081
8.8703 | -88.7087
8.8706 | -88.7097
8.8709 | -88.7105
8.8712 | -88.7112
8.8716 | -88.7117
8.8719 | -88.7124
8.8722 | -88.7137
8.8725 | -88.7144
8.8728 | -88.7153
8.8732 | -88.7165
8.8735 | -88.7169
8.8738 | -88.7179
8.8741 | -88.7188
8.8745 | -88.72
8.8748 | -88.7213
8.8751 | -88.7223
8.8754 | -88.7233
8.8757 | -88.7242
8.8761 | -88.7248
8.8764 | -88.7258
8.8767 | -88.7268
8.877 | -88.7286
8.8773 | -88.7298
8.8777 | -88.7309
8.878 | -88.7313
8.8783 | -88.7318
8.8786 | -88.7327
8.8789 | -88.734
8.8793 | -88.7352
8.8796 | -88.736
8.8799 | -88.7371
8.8802 | -88.7379
8.8806 | -88.7388
8.8809 | -88.7398
8.8812 | -88.7408
8.8815 | -88.7421
8.8818 | -88.7433
8.8822 | -88.7446
8.8825 | -88.7456
8.8828 | -88.7466
8.8831 | -88.7476
8.8834 | -88.7483
8.8838 | -88.7492
8.8841 | -88.7504
8.8844 | -88.7512
8.8847 | -88.7523
8.8851 | -88.7535
8.8854 | -88.7545
8.8857 | -88.756
8.886 | -88.7567
8.8863 | -88.7574
8.8867 | -88.7586
8.887 | -88.7594
8.8873 | -88.7605
8.8876 | -88.7615
8.8879 | -88.7625
8.8883 | -88.7642
8.8886 | -88.7671
8.8889 | -88.7696
8.8892 | -88.7711
8.8896 | -88.7722
8.8899 | -88.7732
8.8902 | -88.7741
8.8905 | -88.7749
8.8908 | -88.776
8.8912 | -88.7771
8.8915 | -88.778
8.8918 | -88.7789
8.8921 | -88.7804
8.8924 | -88.7818
8.8928 | -88.7827
8.8931 | -88.7833
8.8934 | -88.7839
8.8937 | -88.7848
8.8941 | -88.7858
8.8944 | -88.7869
8.8947 | -88.7884
8.895 | -88.7897
8.8953 | -88.7909
8.8957 | -88.7922
8.896 | -88.7935
8.8963 | -88.7939
8.8966 | -88.7949
8.8969 | -88.7957
8.8973 | -88.7962
8.8976 | -88.7974
8.8979 | -88.7989
8.8982 | -88.7995
8.8986 | -88.8003
8.8989 | -88.8013
8.8992 | -88.8019
8.8995 | -88.8029
8.8998 | -88.8037
8.9002 | -88.8048
8.9005 | -88.8058
8.9008 | -88.8067
8.9011 | -88.8076
8.9014 | -88.8085
8.9018 | -88.8092
8.9021 | -88.8102
8.9024 | -88.8118
8.9027 | -88.8136
8.903 | -88.8144
8.9034 | -88.8159
8.9037 | -88.8169
8.904 | -88.8172
8.9043 | -88.8186
8.9047 | -88.8204
8.905 | -88.822
8.9053 | -88.8224
8.9056 | -88.8234
8.9059 | -88.8243
8.9063 | -88.8254
8.9066 | -88.826
8.9069 | -88.8276
8.9072 | -88.8289
8.9075 | -88.83
8.9079 | -88.8309
8.9082 | -88.8313
8.9085 | -88.8322
8.9088 | -88.8333
8.9092 | -88.8346
8.9095 | -88.8356
8.9098 | -88.837
8.9101 | -88.8387
8.9104 | -88.8401
8.9108 | -88.8409
8.9111 | -88.8418
8.9114 | -88.8428
8.9117 | -88.8434
8.912 | -88.8444
8.9124 | -88.8455
8.9127 | -88.8468
8.913 | -88.8481
8.9133 | -88.8494
8.9137 | -88.8501
8.914 | -88.8514
8.9143 | -88.8533
8.9146 | -88.8541
8.9149 | -88.8559
8.9153 | -88.8575
8.9156 | -88.8584
8.9159 | -88.8591
8.9162 | -88.8602
8.9165 | -88.8619
8.9169 | -88.8631
8.9172 | -88.8642
8.9175 | -88.8657
8.9178 | -88.8672
8.9182 | -88.8681
8.9185 | -88.8689
8.9188 | -88.8695
8.9191 | -88.8706
8.9194 | -88.8715
8.9198 | -88.8728
8.9201 | -88.8754
8.9204 | -88.8769
8.9207 | -88.8784
8.921 | -88.8797
8.9214 | -88.8805
8.9217 | -88.8813
8.922 | -88.8822
8.9223 | -88.8836
8.9227 | -88.8859
8.923 | -88.8872
8.9233 | -88.8876
8.9236 | -88.8881
8.9239 | -88.8889
8.9243 | -88.8899
8.9246 | -88.8915
8.9249 | -88.8925
8.9252 | -88.8934
8.9255 | -88.895
8.9259 | -88.897
8.9262 | -88.8989
8.9265 | -88.9007
8.9268 | -88.902
8.9271 | -88.9031
8.9275 | -88.9044
8.9278 | -88.9054
8.9281 | -88.9061
8.9284 | -88.9071
8.9288 | -88.9081
8.9291 | -88.9089
8.9294 | -88.9108
8.9297 | -88.9134
8.93 | -88.9155
8.9304 | -88.917
8.9307 | -88.9182
8.931 | -88.9189
8.9313 | -88.9202
8.9316 | -88.9217
8.932 | -88.9228
8.9323 | -88.9242
8.9326 | -88.9258
8.9329 | -88.9275
8.9333 | -88.9284
8.9336 | -88.9292
8.9339 | -88.9299
8.9342 | -88.9307
8.9345 | -88.9315
8.9349 | -88.9323
8.9352 | -88.9334
8.9355 | -88.9343
8.9358 | -88.9353
8.9361 | -88.9362
8.9365 | -88.9366
8.9368 | -88.9367
8.9371 | -88.9379
8.9374 | -88.9388
8.9378 | -88.9395
8.9381 | -88.9404
8.9384 | -88.9417
8.9387 | -88.9424
8.939 | -88.943
8.9394 | -88.9438
8.9397 | -88.9448
8.94 | -88.946
8.9403 | -88.9467
8.9406 | -88.9476
8.941 | -88.9486
8.9413 | -88.9497
8.9416 | -88.9507
8.9419 | -88.9517
8.9423 | -88.9526
8.9426 | -88.9538
8.9429 | -88.9552
8.9432 | -88.9561
8.9435 | -88.9567
8.9439 | -88.9576
8.9442 | -88.9586
8.9445 | -88.9598
8.9448 | -88.9605
8.9451 | -88.9614
8.9455 | -88.9621
8.9458 | -88.9628
8.9461 | -88.9633
8.9464 | -88.9645
8.9467 | -88.9659
8.9471 | -88.9669
8.9474 | -88.9678
8.9477 | -88.969
8.948 | -88.9702
8.9484 | -88.9712
8.9487 | -88.9723
8.949 | -88.9734
8.9493 | -88.9744
8.9496 | -88.9752
8.95 | -88.9764
8.9503 | -88.9772
8.9506 | -88.9775
8.9509 | -88.9784
8.9512 | -88.9784
8.9516 | -88.9782
8.9519 | -88.9778
8.9522 | -88.9785
8.9525 | -88.9797
8.9529 | -88.9806
8.9532 | -88.9812
8.9535 | -88.982
8.9538 | -88.9828
8.9541 | -88.9837
8.9545 | -88.9847
8.9548 | -88.986
8.9551 | -88.987
8.9554 | -88.9881
8.9557 | -88.9893
8.9561 | -88.9905
8.9564 | -88.9911
8.9567 | -88.9919
8.957 | -88.9927
8.9574 | -88.9934
8.9577 | -88.9953
8.958 | -88.9974
8.9583 | -88.9992
8.9586 | -89.0013
8.959 | -89.0029
8.9593 | -89.0048
8.9596 | -89.0064
8.9599 | -89.0076
8.9602 | -89.0091
8.9606 | -89.0103
8.9609 | -89.0112
8.9612 | -89.012
8.9615 | -89.0139
8.9619 | -89.0159
8.9622 | -89.0172
8.9625 | -89.0182
8.9628 | -89.0193
8.9631 | -89.0202
8.9635 | -89.021
8.9638 | -89.022
8.9641 | -89.0229
8.9644 | -89.0246
8.9647 | -89.0255
8.9651 | -89.0264
8.9654 | -89.0277
8.9657 | -89.0284
8.966 | -89.0293
8.9664 | -89.0303
8.9667 | -89.0313
8.967 | -89.0323
8.9673 | -89.034
8.9676 | -89.035
8.968 | -89.0359
8.9683 | -89.0369
8.9686 | -89.0382
8.9689 | -89.0392
8.9692 | -89.0399
8.9696 | -89.041
8.9699 | -89.042
8.9702 | -89.0431
8.9705 | -89.0443
8.9708 | -89.0451
8.9712 | -89.0464
8.9715 | -89.0474
8.9718 | -89.0482
8.9721 | -89.0491
8.9725 | -89.0499
8.9728 | -89.0507
8.9731 | -89.0517
8.9734 | -89.0527
8.9737 | -89.0537
8.9741 | -89.0546
8.9744 | -89.0557
8.9747 | -89.0567
8.975 | -89.058
8.9753 | -89.059
8.9757 | -89.0599
8.976 | -89.0605
8.9763 | -89.0613
8.9766 | -89.0623
8.977 | -89.0633
8.9773 | -89.0645
8.9776 | -89.0647
8.9779 | -89.0656
8.9782 | -89.0667
8.9786 | -89.0683
8.9789 | -89.0698
8.9792 | -89.0708
8.9795 | -89.0723
8.9798 | -89.0732
8.9802 | -89.0743
8.9805 | -89.0754
8.9808 | -89.0763
8.9811 | -89.0772
8.9815 | -89.0783
8.9818 | -89.0792
8.9821 | -89.0809
8.9824 | -89.0819
8.9827 | -89.0827
8.9831 | -89.0846
8.9834 | -89.0862
8.9837 | -89.0872
8.984 | -89.0881
8.9843 | -89.089
8.9847 | -89.0897
8.985 | -89.0908
8.9853 | -89.0921
8.9856 | -89.093
8.986 | -89.0943
8.9863 | -89.0961
8.9866 | -89.097
8.9869 | -89.0985
8.9872 | -89.0997
8.9876 | -89.1003
8.9879 | -89.1006
8.9882 | -89.1013
8.9885 | -89.1027
8.9888 | -89.1037
8.9892 | -89.1044
8.9895 | -89.1053
8.9898 | -89.1069
8.9901 | -89.1088
8.9904 | -89.1105
8.9908 | -89.1115
8.9911 | -89.1124
8.9914 | -89.1133
8.9917 | -89.1144
8.9921 | -89.116
8.9924 | -89.1174
8.9927 | -89.118
8.993 | -89.119
8.9933 | -89.1205
8.9937 | -89.1218
8.994 | -89.1225
8.9943 | -89.1233
8.9946 | -89.1239
8.9949 | -89.1257
8.9953 | -89.1272
8.9956 | -89.1285
8.9959 | -89.1298
8.9962 | -89.1311
8.9966 | -89.1319
8.9969 | -89.1324
8.9972 | -89.1326
8.9975 | -89.1333
8.9978 | -89.1345
8.9982 | -89.1355
8.9985 | -89.1364
8.9988 | -89.1377
8.9991 | -89.1393
8.9994 | -89.1405
8.9998 | -89.1414
9.0001 | -89.1425
9.0004 | -89.1438
9.0007 | -89.1448
9.0011 | -89.146
9.0014 | -89.1469
9.0017 | -89.1475
9.002 | -89.148
9.0023 | -89.149
9.0027 | -89.1499
9.003 | -89.1509
9.0033 | -89.1522
9.0036 | -89.1535
9.0039 | -89.1549
9.0043 | -89.1561
9.0046 | -89.1572
9.0049 | -89.1581
9.0052 | -89.1589
9.0056 | -89.1599
9.0059 | -89.161
9.0062 | -89.1617
9.0065 | -89.1628
9.0068 | -89.1636
9.0072 | -89.1643
9.0075 | -89.1651
9.0078 | -89.1661
9.0081 | -89.1669
9.0084 | -89.1672
9.0088 | -89.1684
9.0091 | -89.1694
9.0094 | -89.1705
9.0097 | -89.1721
9.0101 | -89.1734
9.0104 | -89.1745
9.0107 | -89.1755
9.011 | -89.1764
9.0113 | -89.1778
9.0117 | -89.1789
9.012 | -89.1799
9.0123 | -89.1807
9.0126 | -89.1817
9.0129 | -89.1825
9.0133 | -89.184
9.0136 | -89.1862
9.0139 | -89.1874
9.0142 | -89.1885
9.0145 | -89.1895
9.0149 | -89.1913
9.0152 | -89.1931
9.0155 | -89.1942
9.0158 | -89.1951
9.0162 | -89.1956
9.0165 | -89.1963
9.0168 | -89.1975
9.0171 | -89.1983
9.0174 | -89.199
9.0178 | -89.2008
9.0181 | -89.2018
9.0184 | -89.203
9.0187 | -89.2047
9.019 | -89.2058
9.0194 | -89.2069
9.0197 | -89.208
9.02 | -89.2087
9.0203 | -89.2099
9.0207 | -89.21
9.021 | -89.2101
9.0213 | -89.2105
9.0216 | -89.2108
9.0219 | -89.2117
9.0223 | -89.2129
9.0226 | -89.2144
9.0229 | -89.2158
9.0232 | -89.2179
9.0235 | -89.2194
9.0239 | -89.2208
9.0242 | -89.2222
9.0245 | -89.2233
9.0248 | -89.2243
9.0252 | -89.2252
9.0255 | -89.2262
9.0258 | -89.2272
9.0261 | -89.2285
9.0264 | -89.2296
9.0268 | -89.23
9.0271 | -89.2307
9.0274 | -89.232
9.0277 | -89.2337
9.028 | -89.235
9.0284 | -89.236
9.0287 | -89.2372
9.029 | -89.2384
9.0293 | -89.2392
9.0297 | -89.2398
9.03 | -89.2404
9.0303 | -89.2413
9.0306 | -89.2427
9.0309 | -89.2443
9.0313 | -89.2449
9.0316 | -89.2457
9.0319 | -89.2471
9.0322 | -89.2477
9.0325 | -89.2475
9.0329 | -89.2477
9.0332 | -89.2487
9.0335 | -89.2497
9.0338 | -89.2509
9.0342 | -89.2522
9.0345 | -89.2531
9.0348 | -89.2546
9.0351 | -89.2558
9.0354 | -89.2567
9.0358 | -89.2579
9.0361 | -89.2594
9.0364 | -89.2604
9.0367 | -89.2613
9.037 | -89.2619
9.0374 | -89.2624
9.0377 | -89.2634
9.038 | -89.2645
9.0383 | -89.2652
9.0386 | -89.2658
9.039 | -89.2666
9.0393 | -89.2674
9.0396 | -89.2681
9.0399 | -89.269
9.0403 | -89.2698
9.0406 | -89.2706
9.0409 | -89.271
9.0412 | -89.2714
9.0415 | -89.2725
9.0419 | -89.2734
9.0422 | -89.2744
9.0425 | -89.2752
9.0428 | -89.2759
9.0431 | -89.2769
9.0435 | -89.2789
9.0438 | -89.2804
9.0441 | -89.2814
9.0444 | -89.2827
9.0448 | -89.2834
9.0451 | -89.2844
9.0454 | -89.2854
9.0457 | -89.2868
9.046 | -89.2881
9.0464 | -89.2894
9.0467 | -89.2905
9.047 | -89.2916
9.0473 | -89.2933
9.0476 | -89.2948
9.048 | -89.2959
9.0483 | -89.2971
9.0486 | -89.2981
9.0489 | -89.2989
9.0493 | -89.2997
9.0496 | -89.3011
9.0499 | -89.3034
9.0502 | -89.3048
9.0505 | -89.3057
9.0509 | -89.3067
9.0512 | -89.3085
9.0515 | -89.3102
9.0518 | -89.3112
9.0521 | -89.3123
9.0525 | -89.3132
9.0528 | -89.3139
9.0531 | -89.3155
9.0534 | -89.3165
9.0538 | -89.3173
9.0541 | -89.3188
9.0544 | -89.3206
9.0547 | -89.322
9.055 | -89.3231
9.0554 | -89.3242
9.0557 | -89.3249
9.056 | -89.326
9.0563 | -89.3271
9.0566 | -89.3286
9.057 | -89.3299
9.0573 | -89.3309
9.0576 | -89.3318
9.0579 | -89.3327
9.0582 | -89.3333
9.0586 | -89.334
9.0589 | -89.3351
9.0592 | -89.3363
9.0595 | -89.3375
9.0599 | -89.3389
9.0602 | -89.34
9.0605 | -89.341
9.0608 | -89.3421
9.0611 | -89.3432
9.0615 | -89.3442
9.0618 | -89.3449
9.0621 | -89.3466
9.0624 | -89.3479
9.0627 | -89.349
9.0631 | -89.35
9.0634 | -89.3509
9.0637 | -89.3518
9.064 | -89.3528
9.0644 | -89.354
9.0647 | -89.3551
9.065 | -89.3557
9.0653 | -89.357
9.0656 | -89.358
9.066 | -89.359
9.0663 | -89.3603
9.0666 | -89.3617
9.0669 | -89.3631
9.0672 | -89.364
9.0676 | -89.3651
9.0679 | -89.3664
9.0682 | -89.3675
9.0685 | -89.3686
9.0689 | -89.3696
9.0692 | -89.3705
9.0695 | -89.3715
9.0698 | -89.3735
9.0701 | -89.3758
9.0705 | -89.3768
9.0708 | -89.3782
9.0711 | -89.3793
9.0714 | -89.3804
9.0717 | -89.3812
9.0721 | -89.3821
9.0724 | -89.3832
9.0727 | -89.3845
9.073 | -89.385
9.0734 | -89.3861
9.0737 | -89.3871
9.074 | -89.3885
9.0743 | -89.3909
9.0746 | -89.3943
9.075 | -89.3957
9.0753 | -89.3969
9.0756 | -89.398
9.0759 | -89.3995
9.0762 | -89.4006
9.0766 | -89.4013
9.0769 | -89.4021
9.0772 | -89.4035
9.0775 | -89.4045
9.0779 | -89.4061
9.0782 | -89.407
9.0785 | -89.4081
9.0788 | -89.4095
9.0791 | -89.4111
9.0795 | -89.4122
9.0798 | -89.413
9.0801 | -89.4154
9.0804 | -89.4174
9.0807 | -89.4184
9.0811 | -89.4191
9.0814 | -89.42
9.0817 | -89.4212
9.082 | -89.4231
9.0823 | -89.4252
9.0827 | -89.4269
9.083 | -89.4279
9.0833 | -89.4288
9.0836 | -89.4296
9.084 | -89.4304
9.0843 | -89.4321
9.0846 | -89.4336
9.0849 | -89.4348
9.0852 | -89.4357
9.0856 | -89.4368
9.0859 | -89.4374
9.0862 | -89.4382
9.0865 | -89.4396
9.0868 | -89.4407
9.0872 | -89.4422
9.0875 | -89.4434
9.0878 | -89.4438
9.0881 | -89.4446
9.0885 | -89.4458
9.0888 | -89.4471
9.0891 | -89.4481
9.0894 | -89.4491
9.0897 | -89.4503
9.0901 | -89.4513
9.0904 | -89.4523
9.0907 | -89.4533
9.091 | -89.4543
9.0913 | -89.4555
9.0917 | -89.4572
9.092 | -89.4581
9.0923 | -89.4593
9.0926 | -89.4604
9.093 | -89.4613
9.0933 | -89.4621
9.0936 | -89.4631
9.0939 | -89.4638
9.0942 | -89.465
9.0946 | -89.4656
9.0949 | -89.4664
9.0952 | -89.4674
9.0955 | -89.4687
9.0958 | -89.4698
9.0962 | -89.4706
9.0965 | -89.4714
9.0968 | -89.4723
9.0971 | -89.4733
9.0975 | -89.4743
9.0978 | -89.4751
9.0981 | -89.476
9.0984 | -89.4772
9.0987 | -89.4782
9.0991 | -89.4792
9.0994 | -89.4801
9.0997 | -89.4811
9.1 | -89.4819
9.1003 | -89.4827
9.1007 | -89.4837
9.101 | -89.4848
9.1013 | -89.4858
9.1016 | -89.4869
9.102 | -89.4877
9.1023 | -89.4882
9.1026 | -89.4887
9.1029 | -89.4896
9.1032 | -89.4905
9.1036 | -89.4908
9.1039 | -89.4917
9.1042 | -89.4929
9.1045 | -89.494
9.1048 | -89.4946
9.1052 | -89.4952
9.1055 | -89.4954
9.1058 | -89.4962
9.1061 | -89.4973
9.1065 | -89.4985
9.1068 | -89.4995
9.1071 | -89.5002
9.1074 | -89.5011
9.1077 | -89.503
9.1081 | -89.5039
9.1084 | -89.5047
9.1087 | -89.5052
9.109 | -89.5064
9.1093 | -89.5076
9.1097 | -89.5084
9.11 | -89.5093
9.1103 | -89.51
9.1106 | -89.5107
9.1109 | -89.5118
9.1113 | -89.513
9.1116 | -89.5139
9.1119 | -89.5147
9.1122 | -89.5157
9.1126 | -89.5168
9.1129 | -89.5179
9.1132 | -89.5181
9.1135 | -89.5194
9.1138 | -89.5204
9.1142 | -89.5217
9.1145 | -89.5231
9.1148 | -89.5244
9.1151 | -89.5262
9.1154 | -89.527
9.1158 | -89.5278
9.1161 | -89.529
9.1164 | -89.5304
9.1167 | -89.5321
9.1171 | -89.5341
9.1174 | -89.5359
9.1177 | -89.5369
9.118 | -89.5379
9.1183 | -89.5388
9.1187 | -89.5396
9.119 | -89.5404
9.1193 | -89.5411
9.1196 | -89.5416
9.1199 | -89.5425
9.1203 | -89.5434
9.1206 | -89.5447
9.1209 | -89.5457
9.1212 | -89.5466
9.1216 | -89.5476
9.1219 | -89.5484
9.1222 | -89.5489
9.1225 | -89.5502
9.1228 | -89.5517
9.1232 | -89.5529
9.1235 | -89.5541
9.1238 | -89.5545
9.1241 | -89.5552
9.1244 | -89.5562
9.1248 | -89.5572
9.1251 | -89.5582
9.1254 | -89.5592
9.1257 | -89.5601
9.1261 | -89.561
9.1264 | -89.5617
9.1267 | -89.5629
9.127 | -89.5636
9.1273 | -89.5651
9.1277 | -89.5662
9.128 | -89.5671
9.1283 | -89.5681
9.1286 | -89.5689
9.1289 | -89.5699
9.1293 | -89.5707
9.1296 | -89.5718
9.1299 | -89.573
9.1302 | -89.5738
9.1305 | -89.5743
9.1309 | -89.5752
9.1312 | -89.5763
9.1315 | -89.5772
9.1318 | -89.5777
9.1322 | -89.5788
9.1325 | -89.5798
9.1328 | -89.5807
9.1331 | -89.5815
9.1334 | -89.5824
9.1338 | -89.5832
9.1341 | -89.5842
9.1344 | -89.5852
9.1347 | -89.5865
9.135 | -89.5879
9.1354 | -89.5896
9.1357 | -89.5914
9.136 | -89.5933
9.1363 | -89.5952
9.1367 | -89.5966
9.137 | -89.5974
9.1373 | -89.5983
9.1376 | -89.599
9.1379 | -89.5998
9.1383 | -89.6004
9.1386 | -89.6015
9.1389 | -89.6023
9.1392 | -89.6037
9.1395 | -89.6057
9.1399 | -89.6065
9.1402 | -89.6072
9.1405 | -89.6081
9.1408 | -89.6088
9.1412 | -89.6099
9.1415 | -89.6106
9.1418 | -89.6109
9.1421 | -89.6119
9.1424 | -89.6131
9.1428 | -89.6144
9.1431 | -89.6151
9.1434 | -89.6147
9.1437 | -89.6139
9.144 | -89.614
9.1444 | -89.6147
9.1447 | -89.6156
9.145 | -89.6166
9.1453 | -89.6183
9.1457 | -89.6198
9.146 | -89.6209
9.1463 | -89.622
9.1466 | -89.623
9.1469 | -89.6238
9.1473 | -89.6247
9.1476 | -89.6256
9.1479 | -89.6263
9.1482 | -89.6282
9.1485 | -89.6298
9.1489 | -89.6307
9.1492 | -89.6316
9.1495 | -89.6327
9.1498 | -89.6336
9.1502 | -89.6344
9.1505 | -89.6352
9.1508 | -89.6362
9.1511 | -89.6373
9.1514 | -89.6384
9.1518 | -89.6392
9.1521 | -89.64
9.1524 | -89.6402
9.1527 | -89.641
9.153 | -89.6416
9.1534 | -89.6427
9.1537 | -89.6436
9.154 | -89.6449
9.1543 | -89.6466
9.1546 | -89.6478
9.155 | -89.649
9.1553 | -89.6499
9.1556 | -89.6507
9.1559 | -89.6514
9.1563 | -89.6524
9.1566 | -89.6534
9.1569 | -89.6542
9.1572 | -89.6553
9.1575 | -89.6559
9.1579 | -89.6566
9.1582 | -89.6574
9.1585 | -89.6584
9.1588 | -89.6593
9.1591 | -89.6602
9.1595 | -89.661
9.1598 | -89.662
9.1601 | -89.6631
9.1604 | -89.6645
9.1608 | -89.6656
9.1611 | -89.6662
9.1614 | -89.6672
9.1617 | -89.6683
9.162 | -89.669
9.1624 | -89.6702
9.1627 | -89.671
9.163 | -89.6718
9.1633 | -89.6727
9.1636 | -89.6727
9.164 | -89.6729
9.1643 | -89.6741
9.1646 | -89.6748
9.1649 | -89.6755
9.1653 | -89.6761
9.1656 | -89.6772
9.1659 | -89.6785
9.1662 | -89.6793
9.1665 | -89.6798
9.1669 | -89.68
9.1672 | -89.681
9.1675 | -89.6823
9.1678 | -89.6833
9.1681 | -89.684
9.1685 | -89.6851
9.1688 | -89.6864
9.1691 | -89.6864
9.1694 | -89.6869
9.1698 | -89.6876
9.1701 | -89.6886
9.1704 | -89.6894
9.1707 | -89.6904
9.171 | -89.6913
9.1714 | -89.6924
9.1717 | -89.6932
9.172 | -89.694
9.1723 | -89.6949
9.1726 | -89.6963
9.173 | -89.6982
9.1733 | -89.6992
9.1736 | -89.7001
9.1739 | -89.7011
9.1743 | -89.7028
9.1746 | -89.7038
9.1749 | -89.7047
9.1752 | -89.7056
9.1755 | -89.7065
9.1759 | -89.7074
9.1762 | -89.7082
9.1765 | -89.7098
9.1768 | -89.7113
9.1771 | -89.7123
9.1775 | -89.7139
9.1778 | -89.7147
9.1781 | -89.7149
9.1784 | -89.7151
9.1787 | -89.7159
9.1791 | -89.7167
9.1794 | -89.7175
9.1797 | -89.7184
9.18 | -89.7191
9.1804 | -89.7199
9.1807 | -89.7208
9.181 | -89.7213
9.1813 | -89.7221
9.1816 | -89.723
9.182 | -89.7231
9.1823 | -89.7239
9.1826 | -89.7252
9.1829 | -89.7253
9.1832 | -89.7262
9.1836 | -89.7272
9.1839 | -89.728
9.1842 | -89.7287
9.1845 | -89.7292
9.1849 | -89.73
9.1852 | -89.7307
9.1855 | -89.7315
9.1858 | -89.7326
9.1861 | -89.7333
9.1865 | -89.7338
9.1868 | -89.7344
9.1871 | -89.7351
9.1874 | -89.7362
9.1877 | -89.7371
9.1881 | -89.7377
9.1884 | -89.7385
9.1887 | -89.7395
9.189 | -89.7404
9.1894 | -89.7413
9.1897 | -89.7422
9.19 | -89.7429
9.1903 | -89.7431
9.1906 | -89.7436
9.191 | -89.7442
9.1913 | -89.7452
9.1916 | -89.7467
9.1919 | -89.7483
9.1922 | -89.7495
9.1926 | -89.7499
9.1929 | -89.7505
9.1932 | -89.7511
9.1935 | -89.7515
9.1939 | -89.7521
9.1942 | -89.753
9.1945 | -89.7539
9.1948 | -89.7548
9.1951 | -89.7558
9.1955 | -89.757
9.1958 | -89.7575
9.1961 | -89.7584
9.1964 | -89.759
9.1967 | -89.7595
9.1971 | -89.7603
9.1974 | -89.761
9.1977 | -89.7621
9.198 | -89.7628
9.1983 | -89.7635
9.1987 | -89.7647
9.199 | -89.7657
9.1993 | -89.7666
9.1996 | -89.7677
9.2 | -89.7683
9.2003 | -89.7696
9.2006 | -89.7706
9.2009 | -89.7714
9.2012 | -89.7719
9.2016 | -89.7729
9.2019 | -89.7742
9.2022 | -89.7753
9.2025 | -89.7753
9.2028 | -89.7751
9.2032 | -89.7755
9.2035 | -89.7762
9.2038 | -89.777
9.2041 | -89.7778
9.2045 | -89.7786
9.2048 | -89.7794
9.2051 | -89.7802
9.2054 | -89.7809
9.2057 | -89.7815
9.2061 | -89.7824
9.2064 | -89.7828
9.2067 | -89.7834
9.207 | -89.7842
9.2073 | -89.7849
9.2077 | -89.7857
9.208 | -89.7863
9.2083 | -89.7872
9.2086 | -89.7883
9.209 | -89.7892
9.2093 | -89.7897
9.2096 | -89.7905
9.2099 | -89.7916
9.2102 | -89.793
9.2106 | -89.7936
9.2109 | -89.7937
9.2112 | -89.7946
9.2115 | -89.7954
9.2118 | -89.7969
9.2122 | -89.7987
9.2125 | -89.801
9.2128 | -89.8025
9.2131 | -89.8032
9.2135 | -89.8044
9.2138 | -89.8055
9.2141 | -89.8063
9.2144 | -89.8071
9.2147 | -89.808
9.2151 | -89.8089
9.2154 | -89.8088
9.2157 | -89.809
9.216 | -89.8093
9.2163 | -89.8099
9.2167 | -89.8109
9.217 | -89.8117
9.2173 | -89.8125
9.2176 | -89.8132
9.218 | -89.8138
9.2183 | -89.8142
9.2186 | -89.8153
9.2189 | -89.8165
9.2192 | -89.8171
9.2196 | -89.8173
9.2199 | -89.8182
9.2202 | -89.8186
9.2205 | -89.8192
9.2208 | -89.8198
9.2212 | -89.8207
9.2215 | -89.8216
9.2218 | -89.822
9.2221 | -89.8227
9.2224 | -89.8236
9.2228 | -89.8243
9.2231 | -89.825
9.2234 | -89.8255
9.2237 | -89.8262
9.2241 | -89.8275
9.2244 | -89.8282
9.2247 | -89.8289
9.225 | -89.8295
9.2253 | -89.8307
9.2257 | -89.8312
9.226 | -89.8317
9.2263 | -89.8327
9.2266 | -89.8338
9.2269 | -89.8344
9.2273 | -89.8342
9.2276 | -89.8343
9.2279 | -89.8351
9.2282 | -89.8359
9.2286 | -89.8373
9.2289 | -89.8387
9.2292 | -89.8392
9.2295 | -89.8397
9.2298 | -89.8404
9.2302 | -89.8412
9.2305 | -89.8419
9.2308 | -89.8428
9.2311 | -89.8439
9.2314 | -89.8445
9.2318 | -89.8452
9.2321 | -89.846
9.2324 | -89.8464
9.2327 | -89.847
9.2331 | -89.8479
9.2334 | -89.8485
9.2337 | -89.8494
9.234 | -89.8503
9.2343 | -89.8512
9.2347 | -89.8523
9.235 | -89.8532
9.2353 | -89.8538
9.2356 | -89.8545
9.2359 | -89.8543
9.2363 | -89.8548
9.2366 | -89.8555
9.2369 | -89.856
9.2372 | -89.8563
9.2376 | -89.8567
9.2379 | -89.8574
9.2382 | -89.858
9.2385 | -89.8586
9.2388 | -89.8593
9.2392 | -89.8596
9.2395 | -89.8607
9.2398 | -89.8619
9.2401 | -89.8627
9.2404 | -89.8633
9.2408 | -89.8645
9.2411 | -89.8649
9.2414 | -89.8654
9.2417 | -89.8662
9.2421 | -89.867
9.2424 | -89.8674
9.2427 | -89.8673
9.243 | -89.8678
9.2433 | -89.8686
9.2437 | -89.8698
9.244 | -89.8706
9.2443 | -89.8712
9.2446 | -89.8719
9.2449 | -89.8728
9.2453 | -89.8741
9.2456 | -89.8747
9.2459 | -89.8753
9.2462 | -89.8759
9.2465 | -89.8766
9.2469 | -89.8772
9.2472 | -89.878
9.2475 | -89.8781
9.2478 | -89.8783
9.2482 | -89.8788
9.2485 | -89.879
9.2488 | -89.8795
9.2491 | -89.88
9.2494 | -89.8807
9.2498 | -89.8814
9.2501 | -89.8819
9.2504 | -89.8823
9.2507 | -89.8831
9.251 | -89.8844
9.2514 | -89.8855
9.2517 | -89.8867
9.252 | -89.8876
9.2523 | -89.8881
9.2527 | -89.8885
9.253 | -89.889
9.2533 | -89.8889
9.2536 | -89.889
9.2539 | -89.8896
9.2543 | -89.8908
9.2546 | -89.8928
9.2549 | -89.8937
9.2552 | -89.8942
9.2555 | -89.8948
9.2559 | -89.8956
9.2562 | -89.8961
9.2565 | -89.8967
9.2568 | -89.8972
9.2572 | -89.8981
9.2575 | -89.8993
9.2578 | -89.9007
9.2581 | -89.9019
9.2584 | -89.9023
9.2588 | -89.9033
9.2591 | -89.904
9.2594 | -89.9043
9.2597 | -89.9056
9.26 | -89.9073
9.2604 | -89.9082
9.2607 | -89.9087
9.261 | -89.909
9.2613 | -89.9096
9.2617 | -89.9095
9.262 | -89.9098
9.2623 | -89.9107
9.2626 | -89.9113
9.2629 | -89.9117
9.2633 | -89.9122
9.2636 | -89.9128
9.2639 | -89.9132
9.2642 | -89.9137
9.2645 | -89.9143
9.2649 | -89.9149
9.2652 | -89.9155
9.2655 | -89.9161
9.2658 | -89.9167
9.2661 | -89.9174
9.2665 | -89.9179
9.2668 | -89.9187
9.2671 | -89.9199
9.2674 | -89.9208
9.2678 | -89.9212
9.2681 | -89.9217
9.2684 | -89.9223
9.2687 | -89.9229
9.269 | -89.9234
9.2694 | -89.9244
9.2697 | -89.9258
9.27 | -89.9269
9.2703 | -89.9275
9.2706 | -89.9284
9.271 | -89.9292
9.2713 | -89.93
9.2716 | -89.9305
9.2719 | -89.931
9.2723 | -89.9316
9.2726 | -89.9323
9.2729 | -89.9334
9.2732 | -89.9344
9.2735 | -89.9346
9.2739 | -89.9356
9.2742 | -89.9365
9.2745 | -89.9363
9.2748 | -89.9366
9.2751 | -89.9372
9.2755 | -89.9379
9.2758 | -89.9388
9.2761 | -89.9393
9.2764 | -89.94
9.2768 | -89.9416
9.2771 | -89.9432
9.2774 | -89.945
9.2777 | -89.9462
9.278 | -89.9466
9.2784 | -89.9471
9.2787 | -89.9476
9.279 | -89.9481
9.2793 | -89.9487
9.2796 | -89.9493
9.28 | -89.9499
9.2803 | -89.9508
9.2806 | -89.9514
9.2809 | -89.9518
9.2813 | -89.9525
9.2816 | -89.9532
9.2819 | -89.9532
9.2822 | -89.9536
9.2825 | -89.9541
9.2829 | -89.9547
9.2832 | -89.9553
9.2835 | -89.9567
9.2838 | -89.9577
9.2841 | -89.9575
9.2845 | -89.9579
9.2848 | -89.9586
9.2851 | -89.9595
9.2854 | -89.9607
9.2858 | -89.9618
9.2861 | -89.9623
9.2864 | -89.9629
9.2867 | -89.9634
9.287 | -89.9643
9.2874 | -89.9649
9.2877 | -89.9655
9.288 | -89.9663
9.2883 | -89.967
9.2886 | -89.9675
9.289 | -89.9678
9.2893 | -89.9682
9.2896 | -89.9688
9.2899 | -89.969
9.2903 | -89.9701
9.2906 | -89.9714
9.2909 | -89.9723
9.2912 | -89.9731
9.2915 | -89.9739
9.2919 | -89.9746
9.2922 | -89.9751
9.2925 | -89.9757
9.2928 | -89.9765
9.2931 | -89.977
9.2935 | -89.9774
9.2938 | -89.9781
9.2941 | -89.9791
9.2944 | -89.98
9.2947 | -89.9808
9.2951 | -89.9815
9.2954 | -89.9819
9.2957 | -89.9825
9.296 | -89.9835
9.2964 | -89.9842
9.2967 | -89.985
9.297 | -89.9857
9.2973 | -89.9863
9.2976 | -89.9872
9.298 | -89.9879
9.2983 | -89.9887
9.2986 | -89.9892
9.2989 | -89.9898
9.2992 | -89.9903
9.2996 | -89.9908
9.2999 | -89.9913
9.3002 | -89.9921
9.3005 | -89.9931
9.3009 | -89.9937
9.3012 | -89.9947
9.3015 | -89.9951
9.3018 | -89.9957
9.3021 | -89.9964
9.3025 | -89.9972
9.3028 | -89.9986
9.3031 | -89.9993
9.3034 | -89.9998
9.3037 | -90.0005
9.3041 | -90.0011
9.3044 | -90.0014
9.3047 | -90.0019
9.305 | -90.0025
9.3054 | -90.0038
9.3057 | -90.0052
9.306 | -90.0059
9.3063 | -90.0065
9.3066 | -90.0063
9.307 | -90.0059
9.3073 | -90.0061
9.3076 | -90.0065
9.3079 | -90.0073
9.3082 | -90.0079
9.3086 | -90.0092
9.3089 | -90.0109
9.3092 | -90.0116
9.3095 | -90.0122
9.3099 | -90.0127
9.3102 | -90.0131
9.3105 | -90.0136
9.3108 | -90.0142
9.3111 | -90.0147
9.3115 | -90.0152
9.3118 | -90.0159
9.3121 | -90.0162
9.3124 | -90.0168
9.3127 | -90.0173
9.3131 | -90.0183
9.3134 | -90.0198
9.3137 | -90.0199
9.314 | -90.0202
9.3143 | -90.0208
9.3147 | -90.0216
9.315 | -90.0221
9.3153 | -90.0228
9.3156 | -90.0234
9.316 | -90.0236
9.3163 | -90.0244
9.3166 | -90.0255
9.3169 | -90.0259
9.3172 | -90.0262
9.3176 | -90.0271
9.3179 | -90.0282
9.3182 | -90.0295
9.3185 | -90.0301
9.3188 | -90.0306
9.3192 | -90.0311
9.3195 | -90.0316
9.3198 | -90.032
9.3201 | -90.0324
9.3205 | -90.0333
9.3208 | -90.0341
9.3211 | -90.0343
9.3214 | -90.0344
9.3217 | -90.0357
9.3221 | -90.0373
9.3224 | -90.0376
9.3227 | -90.0379
9.323 | -90.0389
9.3233 | -90.0396
9.3237 | -90.0403
9.324 | -90.0407
9.3243 | -90.0413
9.3246 | -90.0416
9.325 | -90.0423
9.3253 | -90.0429
9.3256 | -90.0436
9.3259 | -90.0441
9.3262 | -90.0446
9.3266 | -90.0452
9.3269 | -90.0455
9.3272 | -90.0465
9.3275 | -90.0476
9.3278 | -90.0483
9.3282 | -90.0487
9.3285 | -90.0489
9.3288 | -90.0496
9.3291 | -90.0502
9.3295 | -90.0503
9.3298 | -90.05
9.3301 | -90.0506
9.3304 | -90.0514
9.3307 | -90.052
9.3311 | -90.0526
9.3314 | -90.0532
9.3317 | -90.0541
9.332 | -90.055
9.3323 | -90.0555
9.3327 | -90.0557
9.333 | -90.0563
9.3333 | -90.0572
9.3336 | -90.058
9.3339 | -90.0581
9.3343 | -90.0584
9.3346 | -90.0587
9.3349 | -90.0591
9.3352 | -90.0596
9.3356 | -90.06
9.3359 | -90.0605
9.3362 | -90.0609
9.3365 | -90.0614
9.3368 | -90.0618
9.3372 | -90.0625
9.3375 | -90.0634
9.3378 | -90.0639
9.3381 | -90.0641
9.3384 | -90.0646
9.3388 | -90.0654
9.3391 | -90.066
9.3394 | -90.0664
9.3397 | -90.0671
9.3401 | -90.0683
9.3404 | -90.0697
9.3407 | -90.0708
9.341 | -90.0713
9.3413 | -90.0715
9.3417 | -90.071
9.342 | -90.0708
9.3423 | -90.0712
9.3426 | -90.0714
9.3429 | -90.072
9.3433 | -90.0724
9.3436 | -90.0728
9.3439 | -90.0731
9.3442 | -90.0737
9.3446 | -90.0743
9.3449 | -90.0748
9.3452 | -90.0758
9.3455 | -90.0761
9.3458 | -90.0766
9.3462 | -90.0771
9.3465 | -90.0776
9.3468 | -90.0781
9.3471 | -90.0786
9.3474 | -90.0794
9.3478 | -90.0794
9.3481 | -90.0799
9.3484 | -90.0811
9.3487 | -90.0818
9.3491 | -90.0822
9.3494 | -90.0829
9.3497 | -90.0839
9.35 | -90.0844
9.3503 | -90.0851
9.3507 | -90.086
9.351 | -90.0868
9.3513 | -90.0876
9.3516 | -90.0883
9.3519 | -90.0891
9.3523 | -90.0893
9.3526 | -90.0898
9.3529 | -90.0905
9.3532 | -90.0906
9.3536 | -90.0911
9.3539 | -90.0915
9.3542 | -90.0918
9.3545 | -90.0924
9.3548 | -90.093
9.3552 | -90.0935
9.3555 | -90.0943
9.3558 | -90.0948
9.3561 | -90.0954
9.3564 | -90.0962
9.3568 | -90.0966
9.3571 | -90.096
9.3574 | -90.0956
9.3577 | -90.0959
9.3581 | -90.0963
9.3584 | -90.0961
9.3587 | -90.0966
9.359 | -90.0975
9.3593 | -90.098
9.3597 | -90.0985
9.36 | -90.0986
9.3603 | -90.0993
9.3606 | -90.0999
9.3609 | -90.1006
9.3613 | -90.1013
9.3616 | -90.1016
9.3619 | -90.1021
9.3622 | -90.1031
9.3625 | -90.1034
9.3629 | -90.1038
9.3632 | -90.1037
9.3635 | -90.1044
9.3638 | -90.105
9.3642 | -90.1056
9.3645 | -90.1061
9.3648 | -90.1073
9.3651 | -90.1079
9.3654 | -90.1085
9.3658 | -90.1098
9.3661 | -90.1114
9.3664 | -90.1128
9.3667 | -90.1136
9.367 | -90.1142
9.3674 | -90.1149
9.3677 | -90.1154
9.368 | -90.1159
9.3683 | -90.1163
9.3687 | -90.117
9.369 | -90.1184
9.3693 | -90.1199
9.3696 | -90.1209
9.3699 | -90.1214
9.3703 | -90.1215
9.3706 | -90.1217
9.3709 | -90.1216
9.3712 | -90.1215
9.3715 | -90.122
9.3719 | -90.1228
9.3722 | -90.1231
9.3725 | -90.1234
9.3728 | -90.1237
9.3732 | -90.1236
9.3735 | -90.1234
9.3738 | -90.124
9.3741 | -90.1255
9.3744 | -90.1265
9.3748 | -90.1269
9.3751 | -90.1273
9.3754 | -90.1278
9.3757 | -90.1282
9.376 | -90.1288
9.3764 | -90.1291
9.3767 | -90.1293
9.377 | -90.1296
9.3773 | -90.13
9.3777 | -90.1304
9.378 | -90.131
9.3783 | -90.1311
9.3786 | -90.1309
9.3789 | -90.1309
9.3793 | -90.1314
9.3796 | -90.132
9.3799 | -90.1323
9.3802 | -90.1326
9.3805 | -90.133
9.3809 | -90.1335
9.3812 | -90.1339
9.3815 | -90.1346
9.3818 | -90.1357
9.3821 | -90.1359
9.3825 | -90.1363
9.3828 | -90.1372
9.3831 | -90.1378
9.3834 | -90.1383
9.3838 | -90.1388
9.3841 | -90.1392
9.3844 | -90.1396
9.3847 | -90.14
9.385 | -90.1405
9.3854 | -90.1406
9.3857 | -90.1409
9.386 | -90.1419
9.3863 | -90.1431
9.3866 | -90.1435
9.387 | -90.1438
9.3873 | -90.1447
9.3876 | -90.1451
9.3879 | -90.146
9.3883 | -90.1467
9.3886 | -90.147
9.3889 | -90.1476
9.3892 | -90.1479
9.3895 | -90.1479
9.3899 | -90.1477
9.3902 | -90.1482
9.3905 | -90.1485
9.3908 | -90.1495
9.3911 | -90.1505
9.3915 | -90.1499
9.3918 | -90.1501
9.3921 | -90.1507
9.3924 | -90.1507
9.3928 | -90.1506
9.3931 | -90.1508
9.3934 | -90.1497
9.3937 | -90.1502
9.394 | -90.1507
9.3944 | -90.1512
9.3947 | -90.1515
9.395 | -90.1518
9.3953 | -90.1525
9.3956 | -90.1543
9.396 | -90.1553
9.3963 | -90.1557
9.3966 | -90.1558
9.3969 | -90.1562
9.3973 | -90.1567
9.3976 | -90.1573
9.3979 | -90.1576
9.3982 | -90.1584
9.3985 | -90.1582
9.3989 | -90.1588
9.3992 | -90.1593
9.3995 | -90.1598
9.3998 | -90.1603
9.4001 | -90.1613
9.4005 | -90.1618
9.4008 | -90.1619
9.4011 | -90.162
9.4014 | -90.1623
9.4018 | -90.1628
9.4021 | -90.1634
9.4024 | -90.1637
9.4027 | -90.1643
9.403 | -90.1655
9.4034 | -90.1658
9.4037 | -90.166
9.404 | -90.1668
9.4043 | -90.1673
9.4046 | -90.1679
9.405 | -90.1687
9.4053 | -90.1692
9.4056 | -90.1693
9.4059 | -90.1697
9.4062 | -90.1701
9.4066 | -90.1705
9.4069 | -90.171
9.4072 | -90.1716
9.4075 | -90.1719
9.4079 | -90.1722
9.4082 | -90.1726
9.4085 | -90.1732
9.4088 | -90.1737
9.4091 | -90.1743
9.4095 | -90.1747
9.4098 | -90.1752
9.4101 | -90.1759
9.4104 | -90.1775
9.4107 | -90.1782
9.4111 | -90.1783
9.4114 | -90.1783
9.4117 | -90.1786
9.412 | -90.1783
9.4124 | -90.1781
9.4127 | -90.1783
9.413 | -90.1787
9.4133 | -90.179
9.4136 | -90.1792
9.414 | -90.1793
9.4143 | -90.1795
9.4146 | -90.1797
9.4149 | -90.1803
9.4152 | -90.1805
9.4156 | -90.181
9.4159 | -90.181
9.4162 | -90.1813
9.4165 | -90.1819
9.4169 | -90.1824
9.4172 | -90.1826
9.4175 | -90.1832
9.4178 | -90.1837
9.4181 | -90.1841
9.4185 | -90.1846
9.4188 | -90.1847
9.4191 | -90.1853
9.4194 | -90.186
9.4197 | -90.1868
9.4201 | -90.1871
9.4204 | -90.1875
9.4207 | -90.1879
9.421 | -90.1884
9.4214 | -90.1893
9.4217 | -90.1892
9.422 | -90.1889
9.4223 | -90.1892
9.4226 | -90.1894
9.423 | -90.1897
9.4233 | -90.1907
9.4236 | -90.1916
9.4239 | -90.1916
9.4242 | -90.192
9.4246 | -90.1923
9.4249 | -90.1927
9.4252 | -90.1931
9.4255 | -90.1935
9.4259 | -90.1939
9.4262 | -90.1944
9.4265 | -90.1948
9.4268 | -90.1955
9.4271 | -90.1962
9.4275 | -90.1968
9.4278 | -90.1972
9.4281 | -90.1976
9.4284 | -90.1983
9.4287 | -90.1979
9.4291 | -90.1968
9.4294 | -90.1965
9.4297 | -90.1969
9.43 | -90.1976
9.4303 | -90.198
9.4307 | -90.199
9.431 | -90.1992
9.4313 | -90.1993
9.4316 | -90.2001
9.432 | -90.2006
9.4323 | -90.2013
9.4326 | -90.2022
9.4329 | -90.2029
9.4332 | -90.2037
9.4336 | -90.2046
9.4339 | -90.2055
9.4342 | -90.2059
9.4345 | -90.2063
9.4348 | -90.2069
9.4352 | -90.2072
9.4355 | -90.2078
9.4358 | -90.2082
9.4361 | -90.2086
9.4365 | -90.2092
9.4368 | -90.21
9.4371 | -90.2104
9.4374 | -90.2115
9.4377 | -90.2121
9.4381 | -90.2132
9.4384 | -90.2138
9.4387 | -90.2141
9.439 | -90.2145
9.4393 | -90.2148
9.4397 | -90.2151
9.44 | -90.2154
9.4403 | -90.2156
9.4406 | -90.2161
9.441 | -90.2166
9.4413 | -90.2168
9.4416 | -90.2171
9.4419 | -90.218
9.4422 | -90.2188
9.4426 | -90.2192
9.4429 | -90.22
9.4432 | -90.2204
9.4435 | -90.2209
9.4438 | -90.2212
9.4442 | -90.2216
9.4445 | -90.2221
9.4448 | -90.2225
9.4451 | -90.2232
9.4455 | -90.2234
9.4458 | -90.2237
9.4461 | -90.2238
9.4464 | -90.2242
9.4467 | -90.2245
9.4471 | -90.2246
9.4474 | -90.2256
9.4477 | -90.2265
9.448 | -90.2272
9.4483 | -90.2272
9.4487 | -90.2272
9.449 | -90.2275
9.4493 | -90.2289
9.4496 | -90.23
9.4499 | -90.2302
9.4503 | -90.2304
9.4506 | -90.2307
9.4509 | -90.2312
9.4512 | -90.2314
9.4516 | -90.2318
9.4519 | -90.2323
9.4522 | -90.2329
9.4525 | -90.2334
9.4528 | -90.2338
9.4532 | -90.2341
9.4535 | -90.235
9.4538 | -90.2354
9.4541 | -90.2359
9.4544 | -90.2363
9.4548 | -90.2367
9.4551 | -90.2371
9.4554 | -90.2378
9.4557 | -90.2383
9.4561 | -90.2389
9.4564 | -90.2394
9.4567 | -90.2395
9.457 | -90.2388
9.4573 | -90.2377
9.4577 | -90.2378
9.458 | -90.2385
9.4583 | -90.2385
9.4586 | -90.239
9.4589 | -90.2399
9.4593 | -90.2403
9.4596 | -90.2413
9.4599 | -90.2422
9.4602 | -90.243
9.4606 | -90.2439
9.4609 | -90.2443
9.4612 | -90.2451
9.4615 | -90.2461
9.4618 | -90.2469
9.4622 | -90.2477
9.4625 | -90.248
9.4628 | -90.248
9.4631 | -90.2479
9.4634 | -90.2482
9.4638 | -90.2488
9.4641 | -90.2493
9.4644 | -90.2498
9.4647 | -90.2504
9.4651 | -90.2507
9.4654 | -90.2508
9.4657 | -90.2505
9.466 | -90.2511
9.4663 | -90.2518
9.4667 | -90.2523
9.467 | -90.2534
9.4673 | -90.2537
9.4676 | -90.2536
9.4679 | -90.2531
9.4683 | -90.2535
9.4686 | -90.2538
9.4689 | -90.2543
9.4692 | -90.2546
9.4696 | -90.2553
9.4699 | -90.2557
9.4702 | -90.2558
9.4705 | -90.2561
9.4708 | -90.2564
9.4712 | -90.2567
9.4715 | -90.2568
9.4718 | -90.257
9.4721 | -90.2572
9.4724 | -90.2575
9.4728 | -90.2578
9.4731 | -90.2581
9.4734 | -90.2585
9.4737 | -90.2585
9.4741 | -90.2585
9.4744 | -90.2585
9.4747 | -90.2589
9.475 | -90.2596
9.4753 | -90.26
9.4757 | -90.2603
9.476 | -90.2603
9.4763 | -90.26
9.4766 | -90.2601
9.4769 | -90.2604
9.4773 | -90.2607
9.4776 | -90.261
9.4779 | -90.2613
9.4782 | -90.2616
9.4785 | -90.2621
9.4789 | -90.2628
9.4792 | -90.2635
9.4795 | -90.2638
9.4798 | -90.264
9.4802 | -90.2646
9.4805 | -90.265
9.4808 | -90.2651
9.4811 | -90.2655
9.4814 | -90.2661
9.4818 | -90.2664
9.4821 | -90.2673
9.4824 | -90.2681
9.4827 | -90.2684
9.483 | -90.2684
9.4834 | -90.269
9.4837 | -90.2693
9.484 | -90.2692
9.4843 | -90.2694
9.4847 | -90.27
9.485 | -90.2706
9.4853 | -90.2712
9.4856 | -90.272
9.4859 | -90.2726
9.4863 | -90.273
9.4866 | -90.2733
9.4869 | -90.2737
9.4872 | -90.2744
9.4875 | -90.275
9.4879 | -90.2751
9.4882 | -90.2754
9.4885 | -90.2751
9.4888 | -90.2758
9.4892 | -90.2762
9.4895 | -90.2763
9.4898 | -90.2761
9.4901 | -90.2764
9.4904 | -90.2766
9.4908 | -90.2769
9.4911 | -90.2773
9.4914 | -90.2781
9.4917 | -90.2785
9.492 | -90.2785
9.4924 | -90.2788
9.4927 | -90.2792
9.493 | -90.2795
9.4933 | -90.2795
9.4937 | -90.2796
9.494 | -90.2803
9.4943 | -90.2809
9.4946 | -90.2817
9.4949 | -90.2821
9.4953 | -90.2823
9.4956 | -90.2837
9.4959 | -90.2846
9.4962 | -90.285
9.4965 | -90.2854
9.4969 | -90.2863
9.4972 | -90.2867
9.4975 | -90.2871
9.4978 | -90.2873
9.4981 | -90.2874
9.4985 | -90.2878
9.4988 | -90.2881
9.4991 | -90.2886
9.4994 | -90.2891
9.4998 | -90.2897
9.5001 | -90.2899
9.5004 | -90.2902
9.5007 | -90.2909
9.501 | -90.2914
9.5014 | -90.2915
9.5017 | -90.292
9.502 | -90.2923
9.5023 | -90.2924
9.5026 | -90.2921
9.503 | -90.292
9.5033 | -90.2923
9.5036 | -90.2925
9.5039 | -90.2928
9.5043 | -90.2932
9.5046 | -90.294
9.5049 | -90.2946
9.5052 | -90.2958
9.5055 | -90.2962
9.5059 | -90.2973
9.5062 | -90.2985
9.5065 | -90.2998
9.5068 | -90.3004
9.5071 | -90.3005
9.5075 | -90.3009
9.5078 | -90.3017
9.5081 | -90.302
9.5084 | -90.3025
9.5088 | -90.3033
9.5091 | -90.3042
9.5094 | -90.3049
9.5097 | -90.3049
9.51 | -90.3053
9.5104 | -90.3059
9.5107 | -90.3063
9.511 | -90.3066
9.5113 | -90.3079
9.5116 | -90.3083
9.512 | -90.309
9.5123 | -90.3091
9.5126 | -90.3097
9.5129 | -90.3103
9.5133 | -90.3106
9.5136 | -90.3109
9.5139 | -90.3112
9.5142 | -90.3118
9.5145 | -90.3119
9.5149 | -90.3121
9.5152 | -90.3124
9.5155 | -90.3128
9.5158 | -90.3128
9.5161 | -90.3131
9.5165 | -90.3137
9.5168 | -90.3141
9.5171 | -90.3144
9.5174 | -90.3148
9.5177 | -90.3155
9.5181 | -90.3158
9.5184 | -90.3161
9.5187 | -90.3164
9.519 | -90.3165
9.5194 | -90.3162
9.5197 | -90.3161
9.52 | -90.3164
9.5203 | -90.3168
9.5206 | -90.3173
9.521 | -90.3179
9.5213 | -90.3174
9.5216 | -90.3174
9.5219 | -90.3173
9.5222 | -90.318
9.5226 | -90.3184
9.5229 | -90.3188
9.5232 | -90.319
9.5235 | -90.3195
9.5239 | -90.3198
9.5242 | -90.3202
9.5245 | -90.3209
9.5248 | -90.3226
9.5251 | -90.3231
9.5255 | -90.3233
9.5258 | -90.3236
9.5261 | -90.3237
9.5264 | -90.3239
9.5267 | -90.3243
9.5271 | -90.3243
9.5274 | -90.3243
9.5277 | -90.3244
9.528 | -90.3251
9.5284 | -90.3254
9.5287 | -90.3259
9.529 | -90.3263
9.5293 | -90.3264
9.5296 | -90.3266
9.53 | -90.327
9.5303 | -90.3273
9.5306 | -90.328
9.5309 | -90.3284
9.5312 | -90.3287
9.5316 | -90.3288
9.5319 | -90.3293
9.5322 | -90.3295
9.5325 | -90.3301
9.5329 | -90.3312
9.5332 | -90.3314
9.5335 | -90.3315
9.5338 | -90.3316
9.5341 | -90.3309
9.5345 | -90.3306
9.5348 | -90.3307
9.5351 | -90.3303
9.5354 | -90.3304
9.5357 | -90.3308
9.5361 | -90.3307
9.5364 | -90.3312
9.5367 | -90.3319
9.537 | -90.3324
9.5374 | -90.3333
9.5377 | -90.3346
9.538 | -90.335
9.5383 | -90.3354
9.5386 | -90.336
9.539 | -90.336
9.5393 | -90.3359
9.5396 | -90.3358
9.5399 | -90.3363
9.5402 | -90.3368
9.5406 | -90.3368
9.5409 | -90.3371
9.5412 | -90.3375
9.5415 | -90.3379
9.5419 | -90.3384
9.5422 | -90.3383
9.5425 | -90.3386
9.5428 | -90.3389
9.5431 | -90.3396
9.5435 | -90.3398
9.5438 | -90.3406
9.5441 | -90.3423
9.5444 | -90.3427
9.5447 | -90.3431
9.5451 | -90.3438
9.5454 | -90.3447
9.5457 | -90.3447
9.546 | -90.3453
9.5463 | -90.346
9.5467 | -90.3466
9.547 | -90.3465
9.5473 | -90.3465
9.5476 | -90.3465
9.548 | -90.3472
9.5483 | -90.3482
9.5486 | -90.3488
9.5489 | -90.3494
9.5492 | -90.3487
9.5496 | -90.349
9.5499 | -90.3491
9.5502 | -90.3493
9.5505 | -90.3493
9.5508 | -90.3494
9.5512 | -90.3497
9.5515 | -90.3498
9.5518 | -90.3499
9.5521 | -90.3503
9.5525 | -90.3509
9.5528 | -90.3511
9.5531 | -90.3514
9.5534 | -90.3516
9.5537 | -90.3514
9.5541 | -90.3512
9.5544 | -90.3514
9.5547 | -90.3518
9.555 | -90.352
9.5553 | -90.3524
9.5557 | -90.3525
9.556 | -90.3524
9.5563 | -90.3526
9.5566 | -90.3531
9.557 | -90.3533
9.5573 | -90.3536
9.5576 | -90.3537
9.5579 | -90.3538
9.5582 | -90.3543
9.5586 | -90.3546
9.5589 | -90.3548
9.5592 | -90.3553
9.5595 | -90.3565
9.5598 | -90.3571
9.5602 | -90.3575
9.5605 | -90.358
9.5608 | -90.358
9.5611 | -90.3579
9.5615 | -90.358
9.5618 | -90.3583
9.5621 | -90.3586
9.5624 | -90.359
9.5627 | -90.3595
9.5631 | -90.36
9.5634 | -90.3603
9.5637 | -90.3603
9.564 | -90.3607
9.5643 | -90.3607
9.5647 | -90.361
9.565 | -90.3614
9.5653 | -90.3615
9.5656 | -90.3617
9.5659 | -90.3621
9.5663 | -90.3628
9.5666 | -90.3633
9.5669 | -90.3637
9.5672 | -90.364
9.5676 | -90.3643
9.5679 | -90.3649
9.5682 | -90.3654
9.5685 | -90.3651
9.5688 | -90.3646
9.5692 | -90.3644
9.5695 | -90.3642
9.5698 | -90.3643
9.5701 | -90.3646
9.5704 | -90.3644
9.5708 | -90.3647
9.5711 | -90.3648
9.5714 | -90.3648
9.5717 | -90.3652
9.5721 | -90.3654
9.5724 | -90.3653
9.5727 | -90.3651
9.573 | -90.3652
9.5733 | -90.3655
9.5737 | -90.3654
9.574 | -90.3654
9.5743 | -90.3657
9.5746 | -90.3655
9.5749 | -90.3655
9.5753 | -90.3662
9.5756 | -90.3666
9.5759 | -90.3671
9.5762 | -90.3676
9.5766 | -90.3673
9.5769 | -90.366
9.5772 | -90.366
9.5775 | -90.3661
9.5778 | -90.3665
9.5782 | -90.3667
9.5785 | -90.3669
9.5788 | -90.3678
9.5791 | -90.3689
9.5794 | -90.3691
9.5798 | -90.3691
9.5801 | -90.3691
9.5804 | -90.3684
9.5807 | -90.3671
9.5811 | -90.3667
9.5814 | -90.3673
9.5817 | -90.3675
9.582 | -90.3675
9.5823 | -90.3677
9.5827 | -90.3682
9.583 | -90.3687
9.5833 | -90.3692
9.5836 | -90.3697
9.5839 | -90.3699
9.5843 | -90.3702
9.5846 | -90.3704
9.5849 | -90.3705
9.5852 | -90.3706
9.5855 | -90.3707
9.5859 | -90.371
9.5862 | -90.3711
9.5865 | -90.3719
9.5868 | -90.3723
9.5872 | -90.3724
9.5875 | -90.3728
9.5878 | -90.3729
9.5881 | -90.3729
9.5884 | -90.3733
9.5888 | -90.3739
9.5891 | -90.3746
9.5894 | -90.3751
9.5897 | -90.3751
9.59 | -90.375
9.5904 | -90.3749
9.5907 | -90.3744
9.591 | -90.3742
9.5913 | -90.3746
9.5917 | -90.3749
9.592 | -90.3751
9.5923 | -90.3753
9.5926 | -90.3747
9.5929 | -90.3744
9.5933 | -90.3746
9.5936 | -90.375
9.5939 | -90.376
9.5942 | -90.3774
9.5945 | -90.379
9.5949 | -90.3804
9.5952 | -90.3823
9.5955 | -90.382
9.5958 | -90.3824
9.5962 | -90.3826
9.5965 | -90.383
9.5968 | -90.3836
9.5971 | -90.3841
9.5974 | -90.3844
9.5978 | -90.3849
9.5981 | -90.3858
9.5984 | -90.3859
9.5987 | -90.3861
9.599 | -90.3863
9.5994 | -90.3867
9.5997 | -90.3866
9.6 | -90.3863
9.6003 | -90.3864
9.6007 | -90.3863
9.601 | -90.3868
9.6013 | -90.3873
9.6016 | -90.387
9.6019 | -90.3869
9.6023 | -90.387
9.6026 | -90.3876
9.6029 | -90.388
9.6032 | -90.3885
9.6035 | -90.389
9.6039 | -90.3892
9.6042 | -90.3891
9.6045 | -90.3892
9.6048 | -90.3894
9.6052 | -90.3893
9.6055 | -90.3897
9.6058 | -90.3898
9.6061 | -90.3895
9.6064 | -90.3898
9.6068 | -90.39
9.6071 | -90.3901
9.6074 | -90.39
9.6077 | -90.3899
9.608 | -90.3903
9.6084 | -90.3909
9.6087 | -90.3913
9.609 | -90.3916
9.6093 | -90.3923
9.6097 | -90.3923
9.61 | -90.3926
9.6103 | -90.3929
9.6106 | -90.3932
9.6109 | -90.3933
9.6113 | -90.3933
9.6116 | -90.3935
9.6119 | -90.3937
9.6122 | -90.3939
9.6125 | -90.3942
9.6129 | -90.3943
9.6132 | -90.3942
9.6135 | -90.3944
9.6138 | -90.3934
9.6141 | -90.394
9.6145 | -90.3943
9.6148 | -90.3945
9.6151 | -90.395
9.6154 | -90.3947
9.6158 | -90.3946
9.6161 | -90.3949
9.6164 | -90.3959
9.6167 | -90.3963
9.617 | -90.3965
9.6174 | -90.3971
9.6177 | -90.3973
9.618 | -90.3978
9.6183 | -90.3981
9.6186 | -90.3981
9.619 | -90.3984
9.6193 | -90.3989
9.6196 | -90.3992
9.6199 | -90.3993
9.6203 | -90.3992
9.6206 | -90.399
9.6209 | -90.3992
9.6212 | -90.3999
9.6215 | -90.4003
9.6219 | -90.4002
9.6222 | -90.3994
9.6225 | -90.399
9.6228 | -90.3998
9.6231 | -90.4003
9.6235 | -90.4009
9.6238 | -90.401
9.6241 | -90.4012
9.6244 | -90.4013
9.6248 | -90.4013
9.6251 | -90.4023
9.6254 | -90.4028
9.6257 | -90.403
9.626 | -90.4028
9.6264 | -90.4026
9.6267 | -90.4028
9.627 | -90.403
9.6273 | -90.4033
9.6276 | -90.4034
9.628 | -90.4037
9.6283 | -90.4041
9.6286 | -90.4041
9.6289 | -90.4044
9.6293 | -90.4052
9.6296 | -90.4053
9.6299 | -90.4053
9.6302 | -90.4054
9.6305 | -90.4053
9.6309 | -90.4047
9.6312 | -90.4047
9.6315 | -90.4049
9.6318 | -90.4051
9.6321 | -90.4053
9.6325 | -90.4056
9.6328 | -90.4059
9.6331 | -90.4061
9.6334 | -90.4063
9.6337 | -90.4071
9.6341 | -90.4081
9.6344 | -90.4085
9.6347 | -90.4089
9.635 | -90.4094
9.6354 | -90.4101
9.6357 | -90.4103
9.636 | -90.4105
9.6363 | -90.4109
9.6366 | -90.4109
9.637 | -90.411
9.6373 | -90.4109
9.6376 | -90.4113
9.6379 | -90.4116
9.6382 | -90.412
9.6386 | -90.4127
9.6389 | -90.4134
9.6392 | -90.4141
9.6395 | -90.4148
9.6399 | -90.4153
9.6402 | -90.4155
9.6405 | -90.416
9.6408 | -90.4164
9.6411 | -90.4172
9.6415 | -90.4174
9.6418 | -90.4176
9.6421 | -90.4178
9.6424 | -90.4178
9.6427 | -90.4179
9.6431 | -90.4183
9.6434 | -90.4184
9.6437 | -90.4193
9.644 | -90.42
9.6444 | -90.4188
9.6447 | -90.4183
9.645 | -90.4184
9.6453 | -90.4188
9.6456 | -90.4191
9.646 | -90.4194
9.6463 | -90.4195
9.6466 | -90.4197
9.6469 | -90.4198
9.6472 | -90.4196
9.6476 | -90.4191
9.6479 | -90.4191
9.6482 | -90.4196
9.6485 | -90.4202
9.6489 | -90.4206
9.6492 | -90.4213
9.6495 | -90.4213
9.6498 | -90.4209
9.6501 | -90.4212
9.6505 | -90.4214
9.6508 | -90.4217
9.6511 | -90.4219
9.6514 | -90.4223
9.6517 | -90.4223
9.6521 | -90.4218
9.6524 | -90.4216
9.6527 | -90.4219
9.653 | -90.4223
9.6534 | -90.4225
9.6537 | -90.4226
9.654 | -90.4223
9.6543 | -90.4224
9.6546 | -90.4228
9.655 | -90.4229
9.6553 | -90.4228
9.6556 | -90.4233
9.6559 | -90.4244
9.6562 | -90.4245
9.6566 | -90.4245
9.6569 | -90.4247
9.6572 | -90.4247
9.6575 | -90.4247
9.6578 | -90.4245
9.6582 | -90.4241
9.6585 | -90.4241
9.6588 | -90.4244
9.6591 | -90.4244
9.6595 | -90.4237
9.6598 | -90.4231
9.6601 | -90.4234
9.6604 | -90.4226
9.6607 | -90.4227
9.6611 | -90.4229
9.6614 | -90.4232
9.6617 | -90.4233
9.662 | -90.4235
9.6623 | -90.4235
9.6627 | -90.4223
9.663 | -90.4207
9.6633 | -90.42
9.6636 | -90.4196
9.664 | -90.4199
9.6643 | -90.4208
9.6646 | -90.4219
9.6649 | -90.4218
9.6652 | -90.4221
9.6656 | -90.423
9.6659 | -90.4232
9.6662 | -90.4238
9.6665 | -90.4239
9.6668 | -90.424
9.6672 | -90.4241
9.6675 | -90.4242
9.6678 | -90.4241
9.6681 | -90.4238
9.6685 | -90.424
9.6688 | -90.4248
9.6691 | -90.4255
9.6694 | -90.4269
9.6697 | -90.4273
9.6701 | -90.428
9.6704 | -90.4286
9.6707 | -90.4288
9.671 | -90.429
9.6713 | -90.429
9.6717 | -90.4292
9.672 | -90.4291
9.6723 | -90.429
9.6726 | -90.4288
9.673 | -90.429
9.6733 | -90.4288
9.6736 | -90.4289
9.6739 | -90.4294
9.6742 | -90.4298
9.6746 | -90.4297
9.6749 | -90.429
9.6752 | -90.4284
9.6755 | -90.4289
9.6758 | -90.4293
9.6762 | -90.4293
9.6765 | -90.4293
9.6768 | -90.4294
9.6771 | -90.4294
9.6775 | -90.4295
9.6778 | -90.4291
9.6781 | -90.429
9.6784 | -90.4294
9.6787 | -90.4299
9.6791 | -90.431
9.6794 | -90.4309
9.6797 | -90.431
9.68 | -90.431
9.6803 | -90.4307
9.6807 | -90.431
9.681 | -90.4301
9.6813 | -90.4294
9.6816 | -90.43
9.6819 | -90.4309
9.6823 | -90.4312
9.6826 | -90.4311
9.6829 | -90.4309
9.6832 | -90.4309
9.6836 | -90.4311
9.6839 | -90.4315
9.6842 | -90.4316
9.6845 | -90.4313
9.6848 | -90.4315
9.6852 | -90.4317
9.6855 | -90.4319
9.6858 | -90.4319
9.6861 | -90.4319
9.6864 | -90.4327
9.6868 | -90.4335
9.6871 | -90.4336
9.6874 | -90.4334
9.6877 | -90.4339
9.6881 | -90.434
9.6884 | -90.4343
9.6887 | -90.4345
9.689 | -90.4349
9.6893 | -90.4347
9.6897 | -90.4351
9.69 | -90.4351
9.6903 | -90.4353
9.6906 | -90.4356
9.6909 | -90.436
9.6913 | -90.4362
9.6916 | -90.4362
9.6919 | -90.4361
9.6922 | -90.4359
9.6926 | -90.4361
9.6929 | -90.4363
9.6932 | -90.4361
9.6935 | -90.436
9.6938 | -90.4363
9.6942 | -90.4366
9.6945 | -90.4369
9.6948 | -90.4371
9.6951 | -90.4373
9.6954 | -90.4375
9.6958 | -90.4378
9.6961 | -90.4381
9.6964 | -90.4379
9.6967 | -90.4379
9.6971 | -90.438
9.6974 | -90.4381
9.6977 | -90.4382
9.698 | -90.4383
9.6983 | -90.4384
9.6987 | -90.4383
9.699 | -90.4382
9.6993 | -90.438
9.6996 | -90.438
9.6999 | -90.438
9.7003 | -90.4382
9.7006 | -90.4385
9.7009 | -90.4385
9.7012 | -90.4381
9.7015 | -90.4381
9.7019 | -90.4381
9.7022 | -90.4381
9.7025 | -90.4381
9.7028 | -90.4382
9.7032 | -90.4384
9.7035 | -90.438
9.7038 | -90.4373
9.7041 | -90.4375
9.7044 | -90.4376
9.7048 | -90.4376
9.7051 | -90.4372
9.7054 | -90.4366
9.7057 | -90.4367
9.706 | -90.4368
9.7064 | -90.4366
9.7067 | -90.4365
9.707 | -90.4365
9.7073 | -90.4366
9.7077 | -90.4367
9.708 | -90.4369
9.7083 | -90.4371
9.7086 | -90.4373
9.7089 | -90.4376
9.7093 | -90.4376
9.7096 | -90.4377
9.7099 | -90.4379
9.7102 | -90.438
9.7105 | -90.4374
9.7109 | -90.4372
9.7112 | -90.437
9.7115 | -90.4371
9.7118 | -90.4364
9.7122 | -90.4359
9.7125 | -90.436
9.7128 | -90.4365
9.7131 | -90.4372
9.7134 | -90.4373
9.7138 | -90.4373
9.7141 | -90.4373
9.7144 | -90.4372
9.7147 | -90.4374
9.715 | -90.4375
9.7154 | -90.4374
9.7157 | -90.4374
9.716 | -90.4378
9.7163 | -90.4381
9.7167 | -90.4384
9.717 | -90.4384
9.7173 | -90.438
9.7176 | -90.438
9.7179 | -90.4382
9.7183 | -90.4384
9.7186 | -90.4385
9.7189 | -90.4385
9.7192 | -90.4382
9.7195 | -90.4379
9.7199 | -90.438
9.7202 | -90.4381
9.7205 | -90.4382
9.7208 | -90.4385
9.7212 | -90.4386
9.7215 | -90.4386
9.7218 | -90.4382
9.7221 | -90.4381
9.7224 | -90.4378
9.7228 | -90.4363
9.7231 | -90.4355
9.7234 | -90.4353
9.7237 | -90.4348
9.724 | -90.4348
9.7244 | -90.435
9.7247 | -90.4352
9.725 | -90.4355
9.7253 | -90.4357
9.7256 | -90.4355
9.726 | -90.4344
9.7263 | -90.434
9.7266 | -90.4338
9.7269 | -90.4339
9.7273 | -90.4342
9.7276 | -90.4341
9.7279 | -90.4342
9.7282 | -90.4342
9.7285 | -90.4344
9.7289 | -90.4345
9.7292 | -90.4347
9.7295 | -90.4348
9.7298 | -90.4346
9.7301 | -90.4348
9.7305 | -90.4348
9.7308 | -90.4345
9.7311 | -90.4338
9.7314 | -90.4336
9.7318 | -90.4336
9.7321 | -90.4335
9.7324 | -90.4336
9.7327 | -90.4338
9.733 | -90.4337
9.7334 | -90.4339
9.7337 | -90.4345
9.734 | -90.4343
9.7343 | -90.4336
9.7346 | -90.4333
9.735 | -90.4335
9.7353 | -90.4335
9.7356 | -90.4333
9.7359 | -90.4333
9.7363 | -90.4331
9.7366 | -90.4333
9.7369 | -90.4337
9.7372 | -90.4338
9.7375 | -90.4338
9.7379 | -90.4339
9.7382 | -90.4339
9.7385 | -90.434
9.7388 | -90.434
9.7391 | -90.4345
9.7395 | -90.4343
9.7398 | -90.4343
9.7401 | -90.4345
9.7404 | -90.4345
9.7408 | -90.4342
9.7411 | -90.4346
9.7414 | -90.4347
9.7417 | -90.4343
9.742 | -90.4345
9.7424 | -90.4345
9.7427 | -90.4345
9.743 | -90.4345
9.7433 | -90.4341
9.7436 | -90.4342
9.744 | -90.4343
9.7443 | -90.434
9.7446 | -90.4336
9.7449 | -90.4335
9.7453 | -90.4337
9.7456 | -90.4341
9.7459 | -90.4341
9.7462 | -90.4342
9.7465 | -90.4344
9.7469 | -90.4348
9.7472 | -90.435
9.7475 | -90.4353
9.7478 | -90.4354
9.7481 | -90.436
9.7485 | -90.4366
9.7488 | -90.4366
9.7491 | -90.4367
9.7494 | -90.4368
9.7497 | -90.437
9.7501 | -90.4374
9.7504 | -90.4376
9.7507 | -90.438
9.751 | -90.4383
9.7514 | -90.4386
9.7517 | -90.4388
9.752 | -90.4388
9.7523 | -90.4389
9.7526 | -90.4388
9.753 | -90.4373
9.7533 | -90.4373
9.7536 | -90.4373
9.7539 | -90.4373
9.7542 | -90.4377
9.7546 | -90.438
9.7549 | -90.4385
9.7552 | -90.4386
9.7555 | -90.4386
9.7559 | -90.4388
9.7562 | -90.439
9.7565 | -90.4391
9.7568 | -90.4397
9.7571 | -90.4399
9.7575 | -90.4401
9.7578 | -90.441
9.7581 | -90.4412
9.7584 | -90.4413
9.7587 | -90.4413
9.7591 | -90.4411
9.7594 | -90.4409
9.7597 | -90.4407
9.76 | -90.4412
9.7604 | -90.4413
9.7607 | -90.4414
9.761 | -90.4416
9.7613 | -90.442
9.7616 | -90.4426
9.762 | -90.443
9.7623 | -90.4423
9.7626 | -90.4422
9.7629 | -90.4424
9.7632 | -90.4428
9.7636 | -90.4437
9.7639 | -90.4436
9.7642 | -90.4434
9.7645 | -90.4436
9.7649 | -90.4433
9.7652 | -90.4435
9.7655 | -90.4437
9.7658 | -90.4438
9.7661 | -90.4437
9.7665 | -90.4437
9.7668 | -90.4439
9.7671 | -90.4441
9.7674 | -90.4441
9.7677 | -90.4442
9.7681 | -90.4451
9.7684 | -90.4455
9.7687 | -90.4454
9.769 | -90.4455
9.7693 | -90.4451
9.7697 | -90.4449
9.77 | -90.4446
9.7703 | -90.4445
9.7706 | -90.4444
9.771 | -90.4443
9.7713 | -90.4441
9.7716 | -90.4442
9.7719 | -90.4441
9.7722 | -90.444
9.7726 | -90.4436
9.7729 | -90.4432
9.7732 | -90.4431
9.7735 | -90.4431
9.7738 | -90.4432
9.7742 | -90.4432
9.7745 | -90.4427
9.7748 | -90.4432
9.7751 | -90.4437
9.7755 | -90.4438
9.7758 | -90.4439
9.7761 | -90.4439
9.7764 | -90.4439
9.7767 | -90.4439
9.7771 | -90.4437
9.7774 | -90.4438
9.7777 | -90.4443
9.778 | -90.4443
9.7783 | -90.4444
9.7787 | -90.4446
9.779 | -90.4448
9.7793 | -90.4448
9.7796 | -90.4448
9.78 | -90.445
9.7803 | -90.4446
9.7806 | -90.4442
9.7809 | -90.4444
9.7812 | -90.4444
9.7816 | -90.4445
9.7819 | -90.4447
9.7822 | -90.445
9.7825 | -90.4451
9.7828 | -90.4451
9.7832 | -90.4455
9.7835 | -90.4455
9.7838 | -90.4458
9.7841 | -90.4459
9.7845 | -90.4448
9.7848 | -90.4441
9.7851 | -90.4439
9.7854 | -90.4441
9.7857 | -90.4441
9.7861 | -90.444
9.7864 | -90.4442
9.7867 | -90.4447
9.787 | -90.4445
9.7873 | -90.4447
9.7877 | -90.4454
9.788 | -90.446
9.7883 | -90.4464
9.7886 | -90.4465
9.789 | -90.4466
9.7893 | -90.4468
9.7896 | -90.4471
9.7899 | -90.4473
9.7902 | -90.4469
9.7906 | -90.4467
9.7909 | -90.4466
9.7912 | -90.4459
9.7915 | -90.4459
9.7918 | -90.446
9.7922 | -90.4458
9.7925 | -90.4461
9.7928 | -90.4462
9.7931 | -90.4459
9.7934 | -90.4458
9.7938 | -90.4459
9.7941 | -90.4465
9.7944 | -90.4464
9.7947 | -90.4462
9.7951 | -90.4466
9.7954 | -90.4467
9.7957 | -90.4471
9.796 | -90.4477
9.7963 | -90.4477
9.7967 | -90.448
9.797 | -90.4482
9.7973 | -90.4485
9.7976 | -90.4489
9.7979 | -90.4486
9.7983 | -90.4482
9.7986 | -90.4482
9.7989 | -90.4491
9.7992 | -90.4496
9.7996 | -90.4495
9.7999 | -90.4496
9.8002 | -90.4495
9.8005 | -90.45
9.8008 | -90.4516
9.8012 | -90.4532
9.8015 | -90.4537
9.8018 | -90.454
9.8021 | -90.4539
9.8024 | -90.4539
9.8028 | -90.454
9.8031 | -90.4541
9.8034 | -90.4541
9.8037 | -90.4538
9.8041 | -90.4533
9.8044 | -90.4533
9.8047 | -90.4528
9.805 | -90.4527
9.8053 | -90.4532
9.8057 | -90.4532
9.806 | -90.4533
9.8063 | -90.4529
9.8066 | -90.4527
9.8069 | -90.4528
9.8073 | -90.4527
9.8076 | -90.4525
9.8079 | -90.4524
9.8082 | -90.4525
9.8086 | -90.4527
9.8089 | -90.4523
9.8092 | -90.4521
9.8095 | -90.4522
9.8098 | -90.4524
9.8102 | -90.4526
9.8105 | -90.4534
9.8108 | -90.4536
9.8111 | -90.4538
9.8114 | -90.4538
9.8118 | -90.4537
9.8121 | -90.4537
9.8124 | -90.4542
9.8127 | -90.4541
9.8131 | -90.4537
9.8134 | -90.4534
9.8137 | -90.4531
9.814 | -90.4524
9.8143 | -90.452
9.8147 | -90.4512
9.815 | -90.451
9.8153 | -90.4511
9.8156 | -90.4515
9.8159 | -90.4518
9.8163 | -90.4518
9.8166 | -90.4519
9.8169 | -90.4516
9.8172 | -90.4516
9.8176 | -90.4518
9.8179 | -90.4523
9.8182 | -90.4525
9.8185 | -90.4525
9.8188 | -90.4524
9.8192 | -90.4522
9.8195 | -90.4516
9.8198 | -90.4506
9.8201 | -90.451
9.8204 | -90.4517
9.8208 | -90.4523
9.8211 | -90.4525
9.8214 | -90.4528
9.8217 | -90.4528
9.822 | -90.4516
9.8224 | -90.4525
9.8227 | -90.4535
9.823 | -90.4533
9.8233 | -90.4535
9.8237 | -90.454
9.824 | -90.4544
9.8243 | -90.4545
9.8246 | -90.4543
9.8249 | -90.4543
9.8253 | -90.4549
9.8256 | -90.4556
9.8259 | -90.4557
9.8262 | -90.4558
9.8265 | -90.4558
9.8269 | -90.456
9.8272 | -90.4563
9.8275 | -90.4566
9.8278 | -90.4569
9.8282 | -90.457
9.8285 | -90.4578
9.8288 | -90.4583
9.8291 | -90.4585
9.8294 | -90.4583
9.8298 | -90.4583
9.8301 | -90.4583
9.8304 | -90.4583
9.8307 | -90.4584
9.831 | -90.4584
9.8314 | -90.4585
9.8317 | -90.4585
9.832 | -90.4585
9.8323 | -90.4581
9.8327 | -90.4572
9.833 | -90.4569
9.8333 | -90.4568
9.8336 | -90.4563
9.8339 | -90.456
9.8343 | -90.4558
9.8346 | -90.4558
9.8349 | -90.4558
9.8352 | -90.4553
9.8355 | -90.4553
9.8359 | -90.4555
9.8362 | -90.4556
9.8365 | -90.4559
9.8368 | -90.4558
9.8372 | -90.4559
9.8375 | -90.456
9.8378 | -90.4557
9.8381 | -90.4563
9.8384 | -90.4562
9.8388 | -90.4554
9.8391 | -90.4553
9.8394 | -90.4552
9.8397 | -90.4553
9.84 | -90.4554
9.8404 | -90.4556
9.8407 | -90.4557
9.841 | -90.4559
9.8413 | -90.456
9.8416 | -90.4561
9.842 | -90.457
9.8423 | -90.4579
9.8426 | -90.458
9.8429 | -90.4582
9.8433 | -90.4583
9.8436 | -90.4583
9.8439 | -90.458
9.8442 | -90.4576
9.8445 | -90.4578
9.8449 | -90.4574
9.8452 | -90.4573
9.8455 | -90.4572
9.8458 | -90.4573
9.8461 | -90.457
9.8465 | -90.4572
9.8468 | -90.4577
9.8471 | -90.4578
9.8474 | -90.4578
9.8478 | -90.4579
9.8481 | -90.458
9.8484 | -90.4583
9.8487 | -90.4585
9.849 | -90.4586
9.8494 | -90.4582
9.8497 | -90.4577
9.85 | -90.458
9.8503 | -90.458
9.8506 | -90.4583
9.851 | -90.4587
9.8513 | -90.4589
9.8516 | -90.4591
9.8519 | -90.4593
9.8523 | -90.4591
9.8526 | -90.4592
9.8529 | -90.4595
9.8532 | -90.4598
9.8535 | -90.4601
9.8539 | -90.4607
9.8542 | -90.4617
9.8545 | -90.4624
9.8548 | -90.4625
9.8551 | -90.4629
9.8555 | -90.463
9.8558 | -90.4634
9.8561 | -90.4634
9.8564 | -90.4631
9.8568 | -90.4631
9.8571 | -90.4634
9.8574 | -90.4637
9.8577 | -90.464
9.858 | -90.4643
9.8584 | -90.4641
9.8587 | -90.4634
9.859 | -90.4631
9.8593 | -90.4635
9.8596 | -90.4635
9.86 | -90.4632
9.8603 | -90.4631
9.8606 | -90.4633
9.8609 | -90.4633
9.8613 | -90.4635
9.8616 | -90.4635
9.8619 | -90.4634
9.8622 | -90.4636
9.8625 | -90.464
9.8629 | -90.4635
9.8632 | -90.4635
9.8635 | -90.4639
9.8638 | -90.4649
9.8641 | -90.4659
9.8645 | -90.4655
9.8648 | -90.4661
9.8651 | -90.4663
9.8654 | -90.4664
9.8657 | -90.4665
9.8661 | -90.4667
9.8664 | -90.4669
9.8667 | -90.4669
9.867 | -90.4669
9.8674 | -90.4669
9.8677 | -90.4666
9.868 | -90.4667
9.8683 | -90.4671
9.8686 | -90.4672
9.869 | -90.4671
9.8693 | -90.4671
9.8696 | -90.4673
9.8699 | -90.4672
9.8702 | -90.4667
9.8706 | -90.4664
9.8709 | -90.4667
9.8712 | -90.467
9.8715 | -90.4671
9.8719 | -90.4673
9.8722 | -90.4676
9.8725 | -90.4681
9.8728 | -90.4683
9.8731 | -90.4681
9.8735 | -90.4685
9.8738 | -90.4685
9.8741 | -90.4683
9.8744 | -90.4682
9.8747 | -90.4681
9.8751 | -90.4682
9.8754 | -90.4682
9.8757 | -90.468
9.876 | -90.4677
9.8764 | -90.4678
9.8767 | -90.4679
9.877 | -90.4679
9.8773 | -90.4679
9.8776 | -90.4678
9.878 | -90.4678
9.8783 | -90.4685
9.8786 | -90.4695
9.8789 | -90.4698
9.8792 | -90.4702
9.8796 | -90.4707
9.8799 | -90.4705
9.8802 | -90.4706
9.8805 | -90.4708
9.8809 | -90.4716
9.8812 | -90.4717
9.8815 | -90.4718
9.8818 | -90.4721
9.8821 | -90.4723
9.8825 | -90.4721
9.8828 | -90.4722
9.8831 | -90.4724
9.8834 | -90.4724
9.8837 | -90.472
9.8841 | -90.4712
9.8844 | -90.4711
9.8847 | -90.4714
9.885 | -90.4715
9.8854 | -90.4714
9.8857 | -90.4715
9.886 | -90.4715
9.8863 | -90.4716
9.8866 | -90.4715
9.887 | -90.4712
9.8873 | -90.4714
9.8876 | -90.4712
9.8879 | -90.4707
9.8882 | -90.4708
9.8886 | -90.4709
9.8889 | -90.4716
9.8892 | -90.4715
9.8895 | -90.4714
9.8898 | -90.4715
9.8902 | -90.4717
9.8905 | -90.4725
9.8908 | -90.4731
9.8911 | -90.4732
9.8915 | -90.4733
9.8918 | -90.4737
9.8921 | -90.4741
9.8924 | -90.4741
9.8927 | -90.4741
9.8931 | -90.4741
9.8934 | -90.4739
9.8937 | -90.4736
9.894 | -90.4736
9.8943 | -90.4737
9.8947 | -90.4737
9.895 | -90.4732
9.8953 | -90.4733
9.8956 | -90.4734
9.896 | -90.474
9.8963 | -90.4744
9.8966 | -90.4744
9.8969 | -90.4736
9.8972 | -90.4733
9.8976 | -90.4733
9.8979 | -90.4732
9.8982 | -90.473
9.8985 | -90.4732
9.8988 | -90.4728
9.8992 | -90.4728
9.8995 | -90.4729
9.8998 | -90.4735
9.9001 | -90.4735
9.9005 | -90.4737
9.9008 | -90.4736
9.9011 | -90.4737
9.9014 | -90.4733
9.9017 | -90.4731
9.9021 | -90.4729
9.9024 | -90.4727
9.9027 | -90.4727
9.903 | -90.4731
9.9033 | -90.4736
9.9037 | -90.4738
9.904 | -90.4738
9.9043 | -90.4738
9.9046 | -90.4739
9.905 | -90.4742
9.9053 | -90.4752
9.9056 | -90.4753
9.9059 | -90.4756
9.9062 | -90.4761
9.9066 | -90.4763
9.9069 | -90.4769
9.9072 | -90.4774
9.9075 | -90.4775
9.9078 | -90.4774
9.9082 | -90.4774
9.9085 | -90.4777
9.9088 | -90.4777
9.9091 | -90.4778
9.9094 | -90.4778
9.9098 | -90.4779
9.9101 | -90.478
9.9104 | -90.4782
9.9107 | -90.4786
9.9111 | -90.4784
9.9114 | -90.4778
9.9117 | -90.4782
9.912 | -90.479
9.9123 | -90.4789
9.9127 | -90.4786
9.913 | -90.4789
9.9133 | -90.4794
9.9136 | -90.4796
9.9139 | -90.4798
9.9143 | -90.48
9.9146 | -90.4801
9.9149 | -90.4802
9.9152 | -90.4805
9.9156 | -90.4805
9.9159 | -90.4806
9.9162 | -90.4809
9.9165 | -90.4809
9.9168 | -90.4809
9.9172 | -90.4806
9.9175 | -90.4808
9.9178 | -90.4813
9.9181 | -90.4817
9.9184 | -90.4827
9.9188 | -90.4836
9.9191 | -90.4837
9.9194 | -90.483
9.9197 | -90.4826
9.9201 | -90.4831
9.9204 | -90.4835
9.9207 | -90.4837
9.921 | -90.4838
9.9213 | -90.4839
9.9217 | -90.4836
9.922 | -90.4839
9.9223 | -90.4844
9.9226 | -90.4847
9.9229 | -90.4846
9.9233 | -90.4845
9.9236 | -90.4849
9.9239 | -90.4851
9.9242 | -90.485
9.9246 | -90.4851
9.9249 | -90.4854
9.9252 | -90.4855
9.9255 | -90.4858
9.9258 | -90.4859
9.9262 | -90.486
9.9265 | -90.4865
9.9268 | -90.4869
9.9271 | -90.4874
9.9274 | -90.4881
9.9278 | -90.4882
9.9281 | -90.4881
9.9284 | -90.4879
9.9287 | -90.4879
9.9291 | -90.4885
9.9294 | -90.4884
9.9297 | -90.4882
9.93 | -90.4883
9.9303 | -90.4884
9.9307 | -90.4887
9.931 | -90.4891
9.9313 | -90.4892
9.9316 | -90.4893
9.9319 | -90.4893
9.9323 | -90.4894
9.9326 | -90.4891
9.9329 | -90.489
9.9332 | -90.4895
9.9335 | -90.4904
9.9339 | -90.4904
9.9342 | -90.4901
9.9345 | -90.4901
9.9348 | -90.4905
9.9352 | -90.491
9.9355 | -90.4911
9.9358 | -90.4909
9.9361 | -90.491
9.9364 | -90.4914
9.9368 | -90.4917
9.9371 | -90.4918
9.9374 | -90.4916
9.9377 | -90.4917
9.938 | -90.4915
9.9384 | -90.4915
9.9387 | -90.4917
9.939 | -90.4917
9.9393 | -90.4916
9.9397 | -90.4915
9.94 | -90.4914
9.9403 | -90.4921
9.9406 | -90.4924
9.9409 | -90.4922
9.9413 | -90.4926
9.9416 | -90.493
9.9419 | -90.4931
9.9422 | -90.4931
9.9425 | -90.4931
9.9429 | -90.4929
9.9432 | -90.4924
9.9435 | -90.4921
9.9438 | -90.4923
9.9442 | -90.4926
9.9445 | -90.4927
9.9448 | -90.493
9.9451 | -90.493
9.9454 | -90.4931
9.9458 | -90.493
9.9461 | -90.4929
9.9464 | -90.4928
9.9467 | -90.4926
9.947 | -90.4927
9.9474 | -90.4926
9.9477 | -90.4929
9.948 | -90.4927
9.9483 | -90.4927
9.9487 | -90.4927
9.949 | -90.493
9.9493 | -90.4937
9.9496 | -90.4935
9.9499 | -90.4926
9.9503 | -90.4919
9.9506 | -90.492
9.9509 | -90.4922
9.9512 | -90.4924
9.9515 | -90.4938
9.9519 | -90.4951
9.9522 | -90.4956
9.9525 | -90.4956
9.9528 | -90.4951
9.9532 | -90.495
9.9535 | -90.4948
9.9538 | -90.4938
9.9541 | -90.4932
9.9544 | -90.4934
9.9548 | -90.4935
9.9551 | -90.4936
9.9554 | -90.4937
9.9557 | -90.494
9.956 | -90.4943
9.9564 | -90.4942
9.9567 | -90.4942
9.957 | -90.4947
9.9573 | -90.4948
9.9576 | -90.4954
9.958 | -90.496
9.9583 | -90.4963
9.9586 | -90.4962
9.9589 | -90.4964
9.9593 | -90.4965
9.9596 | -90.4967
9.9599 | -90.4967
9.9602 | -90.4969
9.9605 | -90.497
9.9609 | -90.497
9.9612 | -90.4975
9.9615 | -90.4984
9.9618 | -90.4986
9.9621 | -90.4987
9.9625 | -90.4989
9.9628 | -90.4987
9.9631 | -90.4986
9.9634 | -90.4987
9.9638 | -90.4988
9.9641 | -90.499
9.9644 | -90.4988
9.9647 | -90.4988
9.965 | -90.4988
9.9654 | -90.4981
9.9657 | -90.498
9.966 | -90.498
9.9663 | -90.4976
9.9666 | -90.4978
9.967 | -90.498
9.9673 | -90.4981
9.9676 | -90.4982
9.9679 | -90.4983
9.9683 | -90.4984
9.9686 | -90.4983
9.9689 | -90.4978
9.9692 | -90.4971
9.9695 | -90.4967
9.9699 | -90.4972
9.9702 | -90.4977
9.9705 | -90.4974
9.9708 | -90.4969
9.9711 | -90.4971
9.9715 | -90.497
9.9718 | -90.496
9.9721 | -90.4949
9.9724 | -90.4944
9.9728 | -90.493
9.9731 | -90.4932
9.9734 | -90.4932
9.9737 | -90.4929
9.974 | -90.4928
9.9744 | -90.4928
9.9747 | -90.4914
9.975 | -90.4905
9.9753 | -90.4912
9.9756 | -90.4911
9.976 | -90.4919
9.9763 | -90.4921
9.9766 | -90.4924
9.9769 | -90.4923
9.9772 | -90.4921
9.9776 | -90.4924
9.9779 | -90.4935
9.9782 | -90.494
9.9785 | -90.494
9.9789 | -90.4939
9.9792 | -90.4938
9.9795 | -90.494
9.9798 | -90.4942
9.9801 | -90.4943
9.9805 | -90.4947
9.9808 | -90.4949
9.9811 | -90.4945
9.9814 | -90.4949
9.9817 | -90.495
9.9821 | -90.4952
9.9824 | -90.4953
9.9827 | -90.495
9.983 | -90.4955
9.9834 | -90.4959
9.9837 | -90.496
9.984 | -90.496
9.9843 | -90.4959
9.9846 | -90.4956
9.985 | -90.4961
9.9853 | -90.4963
9.9856 | -90.4966
9.9859 | -90.4961
9.9862 | -90.496
9.9866 | -90.4966
9.9869 | -90.4972
9.9872 | -90.4973
9.9875 | -90.4974
9.9879 | -90.4969
9.9882 | -90.4969
9.9885 | -90.4976
9.9888 | -90.4977
9.9891 | -90.4979
9.9895 | -90.4982
9.9898 | -90.4983
9.9901 | -90.4978
9.9904 | -90.4973
9.9907 | -90.4974
9.9911 | -90.4978
9.9914 | -90.4983
9.9917 | -90.4985
9.992 | -90.4988
9.9924 | -90.4995
9.9927 | -90.4994
9.993 | -90.4992
9.9933 | -90.4993
9.9936 | -90.4993
9.994 | -90.4996
9.9943 | -90.4998
9.9946 | -90.4999
9.9949 | -90.4999
9.9952 | -90.4998
9.9956 | -90.4999
9.9959 | -90.4995
9.9962 | -90.4991
9.9965 | -90.4989
9.9969 | -90.4991
9.9972 | -90.4993
9.9975 | -90.4997
9.9978 | -90.5002
9.9981 | -90.5003
9.9985 | -90.5005
9.9988 | -90.5003
9.9991 | -90.4999
9.9994 | -90.5
9.9997 | -90.5001
10.0001 | -90.5006
10.0004 | -90.5017
10.0007 | -90.5021
10.001 | -90.5021
10.0014 | -90.5018
10.0017 | -90.5014
10.002 | -90.5014
10.0023 | -90.5014
10.0026 | -90.5014
10.003 | -90.5011
10.0033 | -90.4996
10.0036 | -90.4995
10.0039 | -90.4999
10.0042 | -90.5
10.0046 | -90.5003
10.0049 | -90.5008
10.0052 | -90.5007
10.0055 | -90.5006
10.0058 | -90.5005
10.0062 | -90.5006
10.0065 | -90.5007
10.0068 | -90.5008
10.0071 | -90.5011
10.0075 | -90.5012
10.0078 | -90.5011
10.0081 | -90.501
10.0084 | -90.5008
10.0087 | -90.5006
10.0091 | -90.5006
10.0094 | -90.5009
10.0097 | -90.5011
10.01 | -90.5011
10.0103 | -90.5009
10.0107 | -90.5009
10.011 | -90.5012
10.0113 | -90.5015
10.0116 | -90.5017
10.012 | -90.5019
10.0123 | -90.502
10.0126 | -90.5025
10.0129 | -90.5028
10.0132 | -90.5039
10.0136 | -90.5047
10.0139 | -90.505
10.0142 | -90.5054
10.0145 | -90.5054
10.0148 | -90.5054
10.0152 | -90.5056
10.0155 | -90.506
10.0158 | -90.5061
10.0161 | -90.5061
10.0165 | -90.5062
10.0168 | -90.5062
10.0171 | -90.5062
10.0174 | -90.5064
10.0177 | -90.5067
10.0181 | -90.5065
10.0184 | -90.5067
10.0187 | -90.5068
10.019 | -90.5068
10.0193 | -90.507
10.0197 | -90.5078
10.02 | -90.5084
10.0203 | -90.5086
10.0206 | -90.5088
10.021 | -90.509
10.0213 | -90.5091
10.0216 | -90.5091
10.0219 | -90.509
10.0222 | -90.5097
10.0226 | -90.5102
10.0229 | -90.5103
10.0232 | -90.5104
10.0235 | -90.5104
10.0238 | -90.5105
10.0242 | -90.5107
10.0245 | -90.5111
10.0248 | -90.5118
10.0251 | -90.512
10.0254 | -90.5125
10.0258 | -90.5126
10.0261 | -90.5126
10.0264 | -90.5128
10.0267 | -90.5131
10.0271 | -90.5132
10.0274 | -90.5134
10.0277 | -90.5134
10.028 | -90.5136
10.0283 | -90.5138
10.0287 | -90.5141
10.029 | -90.5146
10.0293 | -90.5153
10.0296 | -90.5165
10.0299 | -90.5176
10.0303 | -90.5176
10.0306 | -90.5175
10.0309 | -90.5183
10.0312 | -90.5187
10.0316 | -90.519
10.0319 | -90.5193
10.0322 | -90.5197
10.0325 | -90.5198
10.0328 | -90.5201
10.0332 | -90.5202
10.0335 | -90.5203
10.0338 | -90.5205
10.0341 | -90.5206
10.0344 | -90.5207
10.0348 | -90.5208
10.0351 | -90.521
10.0354 | -90.5212
10.0357 | -90.5213
10.0361 | -90.5217
10.0364 | -90.5216
10.0367 | -90.5214
10.037 | -90.5213
10.0373 | -90.521
10.0377 | -90.521
10.038 | -90.5211
10.0383 | -90.5211
10.0386 | -90.5211
10.0389 | -90.5212
10.0393 | -90.521
10.0396 | -90.5207
10.0399 | -90.5207
10.0402 | -90.5206
10.0406 | -90.521
10.0409 | -90.521
10.0412 | -90.5211
10.0415 | -90.5211
10.0418 | -90.5207
10.0422 | -90.5208
10.0425 | -90.5209
10.0428 | -90.521
10.0431 | -90.5211
10.0434 | -90.5212
10.0438 | -90.5211
10.0441 | -90.5215
10.0444 | -90.5221
10.0447 | -90.5228
10.045 | -90.5232
10.0454 | -90.5229
10.0457 | -90.523
10.046 | -90.5233
10.0463 | -90.5237
10.0467 | -90.5238
10.047 | -90.5236
10.0473 | -90.524
10.0476 | -90.5241
10.0479 | -90.5239
10.0483 | -90.5238
10.0486 | -90.5236
10.0489 | -90.5237
10.0492 | -90.5238
10.0495 | -90.5238
10.0499 | -90.5239
10.0502 | -90.5238
10.0505 | -90.5239
10.0508 | -90.5238
10.0512 | -90.5237
10.0515 | -90.5235
10.0518 | -90.5235
10.0521 | -90.5235
10.0524 | -90.5238
10.0528 | -90.5244
10.0531 | -90.5249
10.0534 | -90.5246
10.0537 | -90.5237
10.054 | -90.5235
10.0544 | -90.5234
10.0547 | -90.523
10.055 | -90.5232
10.0553 | -90.5232
10.0557 | -90.5232
10.056 | -90.5233
10.0563 | -90.5237
10.0566 | -90.5238
10.0569 | -90.5228
10.0573 | -90.522
10.0576 | -90.522
10.0579 | -90.5221
10.0582 | -90.5221
10.0585 | -90.5223
10.0589 | -90.5226
10.0592 | -90.5229
10.0595 | -90.523
10.0598 | -90.5227
10.0602 | -90.5229
10.0605 | -90.5231
10.0608 | -90.5224
10.0611 | -90.5223
10.0614 | -90.5223
10.0618 | -90.5224
10.0621 | -90.5224
10.0624 | -90.5227
10.0627 | -90.5226
10.063 | -90.5224
10.0634 | -90.5228
10.0637 | -90.5228
10.064 | -90.5229
10.0643 | -90.5235
10.0647 | -90.524
10.065 | -90.5241
10.0653 | -90.5242
10.0656 | -90.5244
10.0659 | -90.5245
10.0663 | -90.5248
10.0666 | -90.5259
10.0669 | -90.5264
10.0672 | -90.5264
10.0675 | -90.5272
10.0679 | -90.528
10.0682 | -90.5283
10.0685 | -90.5286
10.0688 | -90.5302
10.0692 | -90.5302
10.0695 | -90.5303
10.0698 | -90.5309
10.0701 | -90.5313
10.0704 | -90.5316
10.0708 | -90.5316
10.0711 | -90.5315
10.0714 | -90.5312
10.0717 | -90.5313
10.072 | -90.5314
10.0724 | -90.5317
10.0727 | -90.5316
10.073 | -90.5317
10.0733 | -90.5317
10.0736 | -90.5318
10.074 | -90.5326
10.0743 | -90.5327
10.0746 | -90.5329
10.0749 | -90.5331
10.0753 | -90.5331
10.0756 | -90.5333
10.0759 | -90.5337
10.0762 | -90.5339
10.0765 | -90.5342
10.0769 | -90.5339
10.0772 | -90.5335
10.0775 | -90.5331
10.0778 | -90.533
10.0781 | -90.5331
10.0785 | -90.5334
10.0788 | -90.5337
10.0791 | -90.5338
10.0794 | -90.5338
10.0798 | -90.534
10.0801 | -90.5337
10.0804 | -90.5333
10.0807 | -90.5334
10.081 | -90.5336
10.0814 | -90.5341
10.0817 | -90.5346
10.082 | -90.5348
10.0823 | -90.5346
10.0826 | -90.5337
10.083 | -90.5331
10.0833 | -90.5323
10.0836 | -90.532
10.0839 | -90.5322
10.0843 | -90.5323
10.0846 | -90.5324
10.0849 | -90.5325
10.0852 | -90.5325
10.0855 | -90.5325
10.0859 | -90.5325
10.0862 | -90.5326
10.0865 | -90.5327
10.0868 | -90.5329
10.0871 | -90.5327
10.0875 | -90.5327
10.0878 | -90.5333
10.0881 | -90.5334
10.0884 | -90.5334
10.0888 | -90.5342
10.0891 | -90.5341
10.0894 | -90.5342
10.0897 | -90.5342
10.09 | -90.5343
10.0904 | -90.5337
10.0907 | -90.5336
10.091 | -90.5338
10.0913 | -90.5338
10.0916 | -90.5338
10.092 | -90.534
10.0923 | -90.5342
10.0926 | -90.5345
10.0929 | -90.5345
10.0932 | -90.5341
10.0936 | -90.5338
10.0939 | -90.5339
10.0942 | -90.5339
10.0945 | -90.5338
10.0949 | -90.5339
10.0952 | -90.5341
10.0955 | -90.5342
10.0958 | -90.5345
10.0961 | -90.5345
10.0965 | -90.5343
10.0968 | -90.5338
10.0971 | -90.5338
10.0974 | -90.5335
10.0977 | -90.5333
10.0981 | -90.5334
10.0984 | -90.5329
10.0987 | -90.5329
10.099 | -90.5337
10.0994 | -90.5343
10.0997 | -90.5341
10.1 | -90.5344
10.1003 | -90.5345
10.1006 | -90.5347
10.101 | -90.5349
10.1013 | -90.5353
10.1016 | -90.5356
10.1019 | -90.5359
10.1022 | -90.5359
10.1026 | -90.5359
10.1029 | -90.5361
10.1032 | -90.5361
10.1035 | -90.5361
10.1039 | -90.5362
10.1042 | -90.5359
10.1045 | -90.5358
10.1048 | -90.5359
10.1051 | -90.5361
10.1055 | -90.5361
10.1058 | -90.5359
10.1061 | -90.5358
10.1064 | -90.5354
10.1067 | -90.536
10.1071 | -90.5376
10.1074 | -90.5383
10.1077 | -90.5385
10.108 | -90.5385
10.1084 | -90.5385
10.1087 | -90.5388
10.109 | -90.5395
10.1093 | -90.5399
10.1096 | -90.5397
10.11 | -90.5399
10.1103 | -90.5398
10.1106 | -90.5399
10.1109 | -90.5402
10.1112 | -90.5402
10.1116 | -90.5402
10.1119 | -90.5403
10.1122 | -90.5406
10.1125 | -90.5408
10.1129 | -90.5411
10.1132 | -90.5415
10.1135 | -90.5415
10.1138 | -90.5411
10.1141 | -90.5412
10.1145 | -90.5411
10.1148 | -90.5415
10.1151 | -90.5413
10.1154 | -90.5411
10.1157 | -90.5406
10.1161 | -90.5408
10.1164 | -90.5404
10.1167 | -90.5403
10.117 | -90.5401
10.1173 | -90.5404
10.1177 | -90.5403
10.118 | -90.5404
10.1183 | -90.5405
10.1186 | -90.5405
10.119 | -90.541
10.1193 | -90.5411
10.1196 | -90.5409
10.1199 | -90.5407
10.1202 | -90.5408
10.1206 | -90.5409
10.1209 | -90.5412
10.1212 | -90.5409
10.1215 | -90.5409
10.1218 | -90.5409
10.1222 | -90.541
10.1225 | -90.541
10.1228 | -90.5409
10.1231 | -90.5409
10.1235 | -90.541
10.1238 | -90.5415
10.1241 | -90.5416
10.1244 | -90.5417
10.1247 | -90.541
10.1251 | -90.5401
10.1254 | -90.5401
10.1257 | -90.5401
10.126 | -90.5404
10.1263 | -90.5406
10.1267 | -90.54
10.127 | -90.5395
10.1273 | -90.5391
10.1276 | -90.5393
10.128 | -90.5397
10.1283 | -90.5401
10.1286 | -90.54
10.1289 | -90.54
10.1292 | -90.5401
10.1296 | -90.5398
10.1299 | -90.5393
10.1302 | -90.5393
10.1305 | -90.5394
10.1308 | -90.5391
10.1312 | -90.5385
10.1315 | -90.5384
10.1318 | -90.5384
10.1321 | -90.5385
10.1325 | -90.5387
10.1328 | -90.5387
10.1331 | -90.5384
10.1334 | -90.5389
10.1337 | -90.5389
10.1341 | -90.5389
10.1344 | -90.5382
10.1347 | -90.5376
10.135 | -90.5371
10.1353 | -90.5367
10.1357 | -90.5367
10.136 | -90.5367
10.1363 | -90.5368
10.1366 | -90.5371
10.137 | -90.538
10.1373 | -90.5395
10.1376 | -90.54
10.1379 | -90.5402
10.1382 | -90.5406
10.1386 | -90.5405
10.1389 | -90.5407
10.1392 | -90.5409
10.1395 | -90.5407
10.1398 | -90.5408
10.1402 | -90.5407
10.1405 | -90.5401
10.1408 | -90.5396
10.1411 | -90.5396
10.1414 | -90.5397
10.1418 | -90.5396
10.1421 | -90.5395
10.1424 | -90.5397
10.1427 | -90.54
10.1431 | -90.5405
10.1434 | -90.5411
10.1437 | -90.5416
10.144 | -90.5417
10.1443 | -90.5418
10.1447 | -90.542
10.145 | -90.5419
10.1453 | -90.542
10.1456 | -90.5424
10.1459 | -90.5428
10.1463 | -90.5428
10.1466 | -90.5428
10.1469 | -90.5428
10.1472 | -90.5425
10.1476 | -90.5416
10.1479 | -90.5409
10.1482 | -90.5408
10.1485 | -90.5415
10.1488 | -90.5415
10.1492 | -90.5414
10.1495 | -90.5415
10.1498 | -90.5416
10.1501 | -90.5418
10.1504 | -90.5418
10.1508 | -90.5421
10.1511 | -90.5422
10.1514 | -90.5424
10.1517 | -90.5424
10.1521 | -90.5424
10.1524 | -90.5428
10.1527 | -90.5434
10.153 | -90.5441
10.1533 | -90.5439
10.1537 | -90.5436
10.154 | -90.5434
10.1543 | -90.5435
10.1546 | -90.5438
10.1549 | -90.5441
10.1553 | -90.5443
10.1556 | -90.5435
10.1559 | -90.5428
10.1562 | -90.5428
10.1566 | -90.5429
10.1569 | -90.543
10.1572 | -90.5434
10.1575 | -90.544
10.1578 | -90.5442
10.1582 | -90.5439
10.1585 | -90.5435
10.1588 | -90.5431
10.1591 | -90.5432
10.1594 | -90.5432
10.1598 | -90.5432
10.1601 | -90.5428
10.1604 | -90.542
10.1607 | -90.5417
10.161 | -90.5419
10.1614 | -90.5423
10.1617 | -90.5423
10.162 | -90.5417
10.1623 | -90.5411
10.1627 | -90.541
10.163 | -90.5411
10.1633 | -90.5411
10.1636 | -90.541
10.1639 | -90.5408
10.1643 | -90.5401
10.1646 | -90.5393
10.1649 | -90.5389
10.1652 | -90.5388
10.1655 | -90.5391
10.1659 | -90.5387
10.1662 | -90.5385
10.1665 | -90.5385
10.1668 | -90.5384
10.1672 | -90.5383
10.1675 | -90.5383
10.1678 | -90.5384
10.1681 | -90.5384
10.1684 | -90.5385
10.1688 | -90.5385
10.1691 | -90.5385
10.1694 | -90.5386
10.1697 | -90.5389
10.17 | -90.5392
10.1704 | -90.5393
10.1707 | -90.5397
10.171 | -90.5399
10.1713 | -90.5399
10.1717 | -90.5402
10.172 | -90.5408
10.1723 | -90.5412
10.1726 | -90.5415
10.1729 | -90.5417
10.1733 | -90.5421
10.1736 | -90.5423
10.1739 | -90.5423
10.1742 | -90.5422
10.1745 | -90.5422
10.1749 | -90.5421
10.1752 | -90.5421
10.1755 | -90.5421
10.1758 | -90.5428
10.1762 | -90.5431
10.1765 | -90.543
10.1768 | -90.5429
10.1771 | -90.5427
10.1774 | -90.5425
10.1778 | -90.5425
10.1781 | -90.5428
10.1784 | -90.5429
10.1787 | -90.5429
10.179 | -90.543
10.1794 | -90.5433
10.1797 | -90.5432
10.18 | -90.5434
10.1803 | -90.5432
10.1807 | -90.5433
10.181 | -90.5436
10.1813 | -90.5436
10.1816 | -90.5436
10.1819 | -90.5433
10.1823 | -90.543
10.1826 | -90.5427
10.1829 | -90.5429
10.1832 | -90.5429
10.1835 | -90.543
10.1839 | -90.5433
10.1842 | -90.5434
10.1845 | -90.5433
10.1848 | -90.5434
10.1851 | -90.5436
10.1855 | -90.5434
10.1858 | -90.5441
10.1861 | -90.5442
10.1864 | -90.544
10.1868 | -90.544
10.1871 | -90.5443
10.1874 | -90.5448
10.1877 | -90.5449
10.188 | -90.5449
10.1884 | -90.545
10.1887 | -90.5451
10.189 | -90.545
10.1893 | -90.5445
10.1896 | -90.5442
10.19 | -90.5441
10.1903 | -90.5443
10.1906 | -90.5446
10.1909 | -90.5448
10.1913 | -90.5447
10.1916 | -90.5442
10.1919 | -90.5438
10.1922 | -90.5437
10.1925 | -90.5438
10.1929 | -90.5438
10.1932 | -90.5441
10.1935 | -90.5442
10.1938 | -90.5442
10.1941 | -90.5444
10.1945 | -90.5451
10.1948 | -90.5448
10.1951 | -90.5446
10.1954 | -90.5445
10.1958 | -90.5434
10.1961 | -90.543
10.1964 | -90.5429
10.1967 | -90.5431
10.197 | -90.5431
10.1974 | -90.543
10.1977 | -90.543
10.198 | -90.5435
10.1983 | -90.5439
10.1986 | -90.5443
10.199 | -90.5446
10.1993 | -90.5446
10.1996 | -90.5438
10.1999 | -90.5435
10.2003 | -90.5442
10.2006 | -90.5444
10.2009 | -90.5444
10.2012 | -90.5444
10.2015 | -90.5441
10.2019 | -90.5438
10.2022 | -90.5437
10.2025 | -90.5438
10.2028 | -90.544
10.2031 | -90.5442
10.2035 | -90.5447
10.2038 | -90.5451
10.2041 | -90.5451
10.2044 | -90.5452
10.2048 | -90.545
10.2051 | -90.5451
10.2054 | -90.5457
10.2057 | -90.5461
10.206 | -90.5462
10.2064 | -90.5467
10.2067 | -90.5468
10.207 | -90.5469
10.2073 | -90.547
10.2076 | -90.5469
10.208 | -90.5472
10.2083 | -90.5474
10.2086 | -90.5473
10.2089 | -90.5473
10.2092 | -90.5475
10.2096 | -90.5475
10.2099 | -90.5476
10.2102 | -90.5477
10.2105 | -90.5475
10.2109 | -90.5473
10.2112 | -90.5473
10.2115 | -90.5474
10.2118 | -90.5477
10.2121 | -90.5483
10.2125 | -90.5488
10.2128 | -90.5484
10.2131 | -90.5482
10.2134 | -90.5483
10.2137 | -90.5488
10.2141 | -90.5493
10.2144 | -90.5491
10.2147 | -90.549
10.215 | -90.5486
10.2154 | -90.5481
10.2157 | -90.5479
10.216 | -90.5477
10.2163 | -90.5474
10.2166 | -90.5473
10.217 | -90.5474
10.2173 | -90.5475
10.2176 | -90.5477
10.2179 | -90.5477
10.2182 | -90.5474
10.2186 | -90.5474
10.2189 | -90.5476
10.2192 | -90.548
10.2195 | -90.5489
10.2199 | -90.5492
10.2202 | -90.5494
10.2205 | -90.5504
10.2208 | -90.5507
10.2211 | -90.5513
10.2215 | -90.5518
10.2218 | -90.5518
10.2221 | -90.5517
10.2224 | -90.5519
10.2227 | -90.5518
10.2231 | -90.5517
10.2234 | -90.5518
10.2237 | -90.5519
10.224 | -90.5519
10.2244 | -90.5523
10.2247 | -90.5527
10.225 | -90.5529
10.2253 | -90.5533
10.2256 | -90.553
10.226 | -90.5532
10.2263 | -90.5533
10.2266 | -90.5538
10.2269 | -90.5537
10.2272 | -90.5534
10.2276 | -90.5535
10.2279 | -90.5537
10.2282 | -90.5542
10.2285 | -90.5541
10.2288 | -90.5542
10.2292 | -90.5545
10.2295 | -90.5546
10.2298 | -90.5546
10.2301 | -90.5543
10.2305 | -90.5544
10.2308 | -90.5545
10.2311 | -90.5544
10.2314 | -90.5553
10.2317 | -90.5558
10.2321 | -90.5557
10.2324 | -90.5558
10.2327 | -90.5559
10.233 | -90.5562
10.2333 | -90.5565
10.2337 | -90.5566
10.234 | -90.5564
10.2343 | -90.5565
10.2346 | -90.5566
10.235 | -90.5568
10.2353 | -90.5573
10.2356 | -90.5581
10.2359 | -90.5588
10.2362 | -90.5591
10.2366 | -90.5595
10.2369 | -90.56
10.2372 | -90.5606
10.2375 | -90.5609
10.2378 | -90.5609
10.2382 | -90.5614
10.2385 | -90.5617
10.2388 | -90.5614
10.2391 | -90.5609
10.2395 | -90.5609
10.2398 | -90.561
10.2401 | -90.5612
10.2404 | -90.5617
10.2407 | -90.5624
10.2411 | -90.5628
10.2414 | -90.5629
10.2417 | -90.5629
10.242 | -90.563
10.2423 | -90.5634
10.2427 | -90.564
10.243 | -90.5641
10.2433 | -90.5641
10.2436 | -90.5644
10.244 | -90.5646
10.2443 | -90.5652
10.2446 | -90.5654
10.2449 | -90.5655
10.2452 | -90.5657
10.2456 | -90.5656
10.2459 | -90.5652
10.2462 | -90.5656
10.2465 | -90.5659
10.2468 | -90.5658
10.2472 | -90.5658
10.2475 | -90.5659
10.2478 | -90.5661
10.2481 | -90.566
10.2485 | -90.5663
10.2488 | -90.5664
10.2491 | -90.5665
10.2494 | -90.5666
10.2497 | -90.5667
10.2501 | -90.5661
10.2504 | -90.566
10.2507 | -90.566
10.251 | -90.5664
10.2513 | -90.5661
10.2517 | -90.5666
10.252 | -90.567
10.2523 | -90.5675
10.2526 | -90.5677
10.253 | -90.5676
10.2533 | -90.5677
10.2536 | -90.5679
10.2539 | -90.5677
10.2542 | -90.5674
10.2546 | -90.5674
10.2549 | -90.5674
10.2552 | -90.5678
10.2555 | -90.5678
10.2558 | -90.5678
10.2562 | -90.5684
10.2565 | -90.5688
10.2568 | -90.5695
10.2571 | -90.57
10.2574 | -90.5705
10.2578 | -90.5704
10.2581 | -90.5701
10.2584 | -90.5704
10.2587 | -90.5702
10.2591 | -90.5703
10.2594 | -90.5705
10.2597 | -90.5706
10.26 | -90.5705
10.2603 | -90.5705
10.2607 | -90.5703
10.261 | -90.5706
10.2613 | -90.5707
10.2616 | -90.5711
10.2619 | -90.5714
10.2623 | -90.5714
10.2626 | -90.572
10.2629 | -90.5725
10.2632 | -90.5729
10.2636 | -90.573
10.2639 | -90.5731
10.2642 | -90.5731
10.2645 | -90.5729
10.2648 | -90.5736
10.2652 | -90.5747
10.2655 | -90.5756
10.2658 | -90.5759
10.2661 | -90.576
10.2664 | -90.5756
10.2668 | -90.5758
10.2671 | -90.5759
10.2674 | -90.5761
10.2677 | -90.5763
10.2681 | -90.5764
10.2684 | -90.5767
10.2687 | -90.5766
10.269 | -90.5779
10.2693 | -90.5795
10.2697 | -90.5802
10.27 | -90.5804
10.2703 | -90.581
10.2706 | -90.5816
10.2709 | -90.5821
10.2713 | -90.5823
10.2716 | -90.5823
10.2719 | -90.5822
10.2722 | -90.5824
10.2726 | -90.5823
10.2729 | -90.5823
10.2732 | -90.5824
10.2735 | -90.5826
10.2738 | -90.5826
10.2742 | -90.5826
10.2745 | -90.5828
10.2748 | -90.5829
10.2751 | -90.583
10.2754 | -90.5833
10.2758 | -90.5823
10.2761 | -90.5819
10.2764 | -90.5822
10.2767 | -90.5819
10.277 | -90.5819
10.2774 | -90.5818
10.2777 | -90.5822
10.278 | -90.5821
10.2783 | -90.5823
10.2787 | -90.5818
10.279 | -90.5813
10.2793 | -90.5816
10.2796 | -90.5825
10.2799 | -90.5834
10.2803 | -90.5839
10.2806 | -90.5842
10.2809 | -90.5843
10.2812 | -90.5846
10.2815 | -90.5852
10.2819 | -90.5853
10.2822 | -90.5862
10.2825 | -90.5871
10.2828 | -90.5872
10.2832 | -90.5875
10.2835 | -90.5879
10.2838 | -90.5881
10.2841 | -90.5883
10.2844 | -90.5887
10.2848 | -90.5887
10.2851 | -90.5886
10.2854 | -90.5892
10.2857 | -90.5895
10.286 | -90.5892
10.2864 | -90.5887
10.2867 | -90.5886
10.287 | -90.5887
10.2873 | -90.5886
10.2877 | -90.5885
10.288 | -90.5887
10.2883 | -90.5889
10.2886 | -90.5889
10.2889 | -90.589
10.2893 | -90.5889
10.2896 | -90.5893
10.2899 | -90.5897
10.2902 | -90.5899
10.2905 | -90.5897
10.2909 | -90.5896
10.2912 | -90.5891
10.2915 | -90.5892
10.2918 | -90.5896
10.2922 | -90.5901
10.2925 | -90.59
10.2928 | -90.5899
10.2931 | -90.5901
10.2934 | -90.5902
10.2938 | -90.5902
10.2941 | -90.5894
10.2944 | -90.5889
10.2947 | -90.5885
10.295 | -90.5886
10.2954 | -90.589
10.2957 | -90.5894
10.296 | -90.5895
10.2963 | -90.5898
10.2966 | -90.5905
10.297 | -90.5906
10.2973 | -90.5907
10.2976 | -90.5909
10.2979 | -90.5913
10.2983 | -90.5919
10.2986 | -90.5918
10.2989 | -90.592
10.2992 | -90.5923
10.2995 | -90.5924
10.2999 | -90.5922
10.3002 | -90.5925
10.3005 | -90.593
10.3008 | -90.593
10.3011 | -90.5929
10.3015 | -90.5931
10.3018 | -90.593
10.3021 | -90.5934
10.3024 | -90.5938
10.3028 | -90.5938
10.3031 | -90.5939
10.3034 | -90.5942
10.3037 | -90.5945
10.304 | -90.5946
10.3044 | -90.5944
10.3047 | -90.5953
10.305 | -90.5961
10.3053 | -90.5959
10.3056 | -90.5957
10.306 | -90.596
10.3063 | -90.5963
10.3066 | -90.5969
10.3069 | -90.5974
10.3073 | -90.5978
10.3076 | -90.598
10.3079 | -90.5982
10.3082 | -90.5983
10.3085 | -90.5975
10.3089 | -90.597
10.3092 | -90.5972
10.3095 | -90.5972
10.3098 | -90.5976
10.3101 | -90.5982
10.3105 | -90.5982
10.3108 | -90.5983
10.3111 | -90.5984
10.3114 | -90.5987
10.3118 | -90.5989
10.3121 | -90.5995
10.3124 | -90.5998
10.3127 | -90.6
10.313 | -90.6003
10.3134 | -90.6012
10.3137 | -90.6022
10.314 | -90.6027
10.3143 | -90.6029
10.3146 | -90.6031
10.315 | -90.6032
10.3153 | -90.6034
10.3156 | -90.6036
10.3159 | -90.6037
10.3163 | -90.6037
10.3166 | -90.6037
10.3169 | -90.6038
10.3172 | -90.6036
10.3175 | -90.6035
10.3179 | -90.6041
10.3182 | -90.6046
10.3185 | -90.604
10.3188 | -90.604
10.3191 | -90.604
10.3195 | -90.6041
10.3198 | -90.6043
10.3201 | -90.6041
10.3204 | -90.604
10.3208 | -90.6038
10.3211 | -90.6038
10.3214 | -90.6037
10.3217 | -90.604
10.322 | -90.6045
10.3224 | -90.6049
10.3227 | -90.6056
10.323 | -90.6064
10.3233 | -90.6068
10.3236 | -90.6069
10.324 | -90.6066
10.3243 | -90.6066
10.3246 | -90.6068
10.3249 | -90.6068
10.3252 | -90.6066
10.3256 | -90.6067
10.3259 | -90.607
10.3262 | -90.6074
10.3265 | -90.6078
10.3269 | -90.6082
10.3272 | -90.6085
10.3275 | -90.6088
10.3278 | -90.6095
10.3281 | -90.6098
10.3285 | -90.6102
10.3288 | -90.6104
10.3291 | -90.6108
10.3294 | -90.6109
10.3297 | -90.611
10.3301 | -90.6113
10.3304 | -90.6117
10.3307 | -90.6119
10.331 | -90.6115
10.3314 | -90.6112
10.3317 | -90.6115
10.332 | -90.6117
10.3323 | -90.612
10.3326 | -90.6123
10.333 | -90.6125
10.3333 | -90.6127
10.3336 | -90.6128
10.3339 | -90.613
10.3342 | -90.6135
10.3346 | -90.6139
10.3349 | -90.6141
10.3352 | -90.6142
10.3355 | -90.6145
10.3359 | -90.6149
10.3362 | -90.6152
10.3365 | -90.6159
10.3368 | -90.6162
10.3371 | -90.6169
10.3375 | -90.6176
10.3378 | -90.6196
10.3381 | -90.6206
10.3384 | -90.6206
10.3387 | -90.6203
10.3391 | -90.6205
10.3394 | -90.6208
10.3397 | -90.6211
10.34 | -90.6214
10.3404 | -90.6215
10.3407 | -90.6219
10.341 | -90.6227
10.3413 | -90.6233
10.3416 | -90.6233
10.342 | -90.6229
10.3423 | -90.6231
10.3426 | -90.6233
10.3429 | -90.6237
10.3432 | -90.6245
10.3436 | -90.625
10.3439 | -90.625
10.3442 | -90.6253
10.3445 | -90.6256
10.3448 | -90.6261
10.3452 | -90.6264
10.3455 | -90.6267
10.3458 | -90.6272
10.3461 | -90.6277
10.3465 | -90.6283
10.3468 | -90.6286
10.3471 | -90.6288
10.3474 | -90.629
10.3477 | -90.6295
10.3481 | -90.6301
10.3484 | -90.6307
10.3487 | -90.6307
10.349 | -90.6312
10.3493 | -90.6315
10.3497 | -90.6317
10.35 | -90.6319
10.3503 | -90.6317
10.3506 | -90.6321
10.351 | -90.6323
10.3513 | -90.6327
10.3516 | -90.6331
10.3519 | -90.6336
10.3522 | -90.6343
10.3526 | -90.6345
10.3529 | -90.6346
10.3532 | -90.6346
10.3535 | -90.6351
10.3538 | -90.636
10.3542 | -90.6363
10.3545 | -90.6369
10.3548 | -90.6371
10.3551 | -90.6368
10.3555 | -90.6372
10.3558 | -90.6374
10.3561 | -90.6371
10.3564 | -90.6371
10.3567 | -90.6374
10.3571 | -90.6377
10.3574 | -90.6389
10.3577 | -90.6404
10.358 | -90.641
10.3583 | -90.641
10.3587 | -90.641
10.359 | -90.6411
10.3593 | -90.6413
10.3596 | -90.6418
10.36 | -90.6424
10.3603 | -90.6427
10.3606 | -90.643
10.3609 | -90.6433
10.3612 | -90.6433
10.3616 | -90.6431
10.3619 | -90.6429
10.3622 | -90.6431
10.3625 | -90.6436
10.3628 | -90.6447
10.3632 | -90.6454
10.3635 | -90.6456
10.3638 | -90.6457
10.3641 | -90.6459
10.3645 | -90.6461
10.3648 | -90.6464
10.3651 | -90.6468
10.3654 | -90.647
10.3657 | -90.6479
10.3661 | -90.6484
10.3664 | -90.6487
10.3667 | -90.6489
10.367 | -90.6493
10.3673 | -90.6495
10.3677 | -90.6495
10.368 | -90.6499
10.3683 | -90.6501
10.3686 | -90.6504
10.3689 | -90.6506
10.3693 | -90.6508
10.3696 | -90.651
10.3699 | -90.6514
10.3702 | -90.6521
10.3706 | -90.6533
10.3709 | -90.654
10.3712 | -90.6542
10.3715 | -90.6543
10.3718 | -90.6545
10.3722 | -90.6546
10.3725 | -90.6546
10.3728 | -90.6542
10.3731 | -90.6548
10.3734 | -90.6556
10.3738 | -90.6559
10.3741 | -90.6559
10.3744 | -90.6556
10.3747 | -90.6556
10.3751 | -90.6557
10.3754 | -90.6557
10.3757 | -90.6557
10.376 | -90.6561
10.3763 | -90.6567
10.3767 | -90.6571
10.377 | -90.6587
10.3773 | -90.6599
10.3776 | -90.6603
10.3779 | -90.6612
10.3783 | -90.6623
10.3786 | -90.6627
10.3789 | -90.6633
10.3792 | -90.6635
10.3796 | -90.6636
10.3799 | -90.664
10.3802 | -90.6645
10.3805 | -90.6649
10.3808 | -90.6652
10.3812 | -90.6655
10.3815 | -90.6656
10.3818 | -90.666
10.3821 | -90.6661
10.3824 | -90.6657
10.3828 | -90.6654
10.3831 | -90.6655
10.3834 | -90.6662
10.3837 | -90.6667
10.3841 | -90.6672
10.3844 | -90.6677
10.3847 | -90.668
10.385 | -90.6684
10.3853 | -90.6686
10.3857 | -90.6692
10.386 | -90.6697
10.3863 | -90.6701
10.3866 | -90.6703
10.3869 | -90.6704
10.3873 | -90.6708
10.3876 | -90.6715
10.3879 | -90.6723
10.3882 | -90.6727
10.3886 | -90.6731
10.3889 | -90.6737
10.3892 | -90.6741
10.3895 | -90.6741
10.3898 | -90.674
10.3902 | -90.6742
10.3905 | -90.6745
10.3908 | -90.6747
10.3911 | -90.675
10.3914 | -90.6753
10.3918 | -90.6752
10.3921 | -90.6751
10.3924 | -90.6747
10.3927 | -90.6749
10.393 | -90.6749
10.3934 | -90.6752
10.3937 | -90.6754
10.394 | -90.6757
10.3943 | -90.6759
10.3947 | -90.6761
10.395 | -90.6764
10.3953 | -90.6767
10.3956 | -90.6771
10.3959 | -90.6774
10.3963 | -90.678
10.3966 | -90.6784
10.3969 | -90.6787
10.3972 | -90.6788
10.3975 | -90.6787
10.3979 | -90.679
10.3982 | -90.6794
10.3985 | -90.6801
10.3988 | -90.6805
10.3992 | -90.6808
10.3995 | -90.6812
10.3998 | -90.6816
10.4001 | -90.6819
10.4004 | -90.6822
10.4008 | -90.683
10.4011 | -90.6841
10.4014 | -90.6844
10.4017 | -90.685
10.402 | -90.6854
10.4024 | -90.6855
10.4027 | -90.6859
10.403 | -90.6865
10.4033 | -90.6868
10.4037 | -90.6874
10.404 | -90.6884
10.4043 | -90.6889
10.4046 | -90.6891
10.4049 | -90.6895
10.4053 | -90.6898
10.4056 | -90.6902
10.4059 | -90.6905
10.4062 | -90.6908
10.4065 | -90.6911
10.4069 | -90.6916
10.4072 | -90.692
10.4075 | -90.6926
10.4078 | -90.6932
10.4082 | -90.6935
10.4085 | -90.6944
10.4088 | -90.6948
10.4091 | -90.6945
10.4094 | -90.6947
10.4098 | -90.6957
10.4101 | -90.6965
10.4104 | -90.6968
10.4107 | -90.6971
10.411 | -90.6971
10.4114 | -90.6976
10.4117 | -90.6973
10.412 | -90.6968
10.4123 | -90.6969
10.4126 | -90.6976
10.413 | -90.6981
10.4133 | -90.6983
10.4136 | -90.6985
10.4139 | -90.6988
10.4143 | -90.699
10.4146 | -90.699
10.4149 | -90.699
10.4152 | -90.6992
10.4155 | -90.6998
10.4159 | -90.7001
10.4162 | -90.7
10.4165 | -90.6998
10.4168 | -90.7001
10.4171 | -90.7005
10.4175 | -90.7009
10.4178 | -90.7016
10.4181 | -90.7022
10.4184 | -90.7025
10.4188 | -90.7028
10.4191 | -90.7033
10.4194 | -90.7037
10.4197 | -90.7035
10.42 | -90.7038
10.4204 | -90.7047
10.4207 | -90.705
10.421 | -90.7056
10.4213 | -90.7063
10.4216 | -90.7082
10.422 | -90.7087
10.4223 | -90.7091
10.4226 | -90.7103
10.4229 | -90.7109
10.4233 | -90.7112
10.4236 | -90.7115
10.4239 | -90.712
10.4242 | -90.7122
10.4245 | -90.7123
10.4249 | -90.7132
10.4252 | -90.7147
10.4255 | -90.7155
10.4258 | -90.7159
10.4261 | -90.7164
10.4265 | -90.717
10.4268 | -90.7174
10.4271 | -90.7175
10.4274 | -90.718
10.4278 | -90.7183
10.4281 | -90.7185
10.4284 | -90.7191
10.4287 | -90.7197
10.429 | -90.7201
10.4294 | -90.7203
10.4297 | -90.7203
10.43 | -90.7207
10.4303 | -90.7215
10.4306 | -90.7225
10.431 | -90.7229
10.4313 | -90.7233
10.4316 | -90.7235
10.4319 | -90.7237
10.4323 | -90.724
10.4326 | -90.7244
10.4329 | -90.7248
10.4332 | -90.725
10.4335 | -90.7253
10.4339 | -90.7261
10.4342 | -90.7268
10.4345 | -90.7273
10.4348 | -90.7277
10.4351 | -90.7282
10.4355 | -90.7286
10.4358 | -90.7291
10.4361 | -90.7296
10.4364 | -90.73
10.4367 | -90.7304
10.4371 | -90.7306
10.4374 | -90.7308
10.4377 | -90.7311
10.438 | -90.7315
10.4384 | -90.7316
10.4387 | -90.7321
10.439 | -90.7326
10.4393 | -90.7328
10.4396 | -90.7335
10.44 | -90.734
10.4403 | -90.7344
10.4406 | -90.7348
10.4409 | -90.7353
10.4412 | -90.7358
10.4416 | -90.7364
10.4419 | -90.7372
10.4422 | -90.7377
10.4425 | -90.7381
10.4429 | -90.7384
10.4432 | -90.7388
10.4435 | -90.7395
10.4438 | -90.7399
10.4441 | -90.7402
10.4445 | -90.7405
10.4448 | -90.7412
10.4451 | -90.7419
10.4454 | -90.7418
10.4457 | -90.7421
10.4461 | -90.7423
10.4464 | -90.7427
10.4467 | -90.7431
10.447 | -90.7433
10.4474 | -90.7433
10.4477 | -90.7432
10.448 | -90.7433
10.4483 | -90.7441
10.4486 | -90.7447
10.449 | -90.7452
10.4493 | -90.7456
10.4496 | -90.7461
10.4499 | -90.7464
10.4502 | -90.7467
10.4506 | -90.7471
10.4509 | -90.7475
10.4512 | -90.7479
10.4515 | -90.7483
10.4519 | -90.7488
10.4522 | -90.7492
10.4525 | -90.7494
10.4528 | -90.7497
10.4531 | -90.7499
10.4535 | -90.7502
10.4538 | -90.7515
10.4541 | -90.7529
10.4544 | -90.7529
10.4547 | -90.7533
10.4551 | -90.7537
10.4554 | -90.7539
10.4557 | -90.7542
10.456 | -90.7546
10.4564 | -90.7547
10.4567 | -90.7551
10.457 | -90.7553
10.4573 | -90.7556
10.4576 | -90.7563
10.458 | -90.7564
10.4583 | -90.7565
10.4586 | -90.757
10.4589 | -90.7575
10.4592 | -90.7578
10.4596 | -90.758
10.4599 | -90.7585
10.4602 | -90.7588
10.4605 | -90.7594
10.4608 | -90.7603
10.4612 | -90.7607
10.4615 | -90.761
10.4618 | -90.7614
10.4621 | -90.7618
10.4625 | -90.7622
10.4628 | -90.7624
10.4631 | -90.7626
10.4634 | -90.763
10.4637 | -90.7636
10.4641 | -90.7646
10.4644 | -90.7651
10.4647 | -90.7654
10.465 | -90.7657
10.4653 | -90.7662
10.4657 | -90.7666
10.466 | -90.7671
10.4663 | -90.7676
10.4666 | -90.7682
10.467 | -90.7689
10.4673 | -90.7695
10.4676 | -90.7698
10.4679 | -90.7704
10.4682 | -90.771
10.4686 | -90.7716
10.4689 | -90.7718
10.4692 | -90.772
10.4695 | -90.7725
10.4698 | -90.7732
10.4702 | -90.7734
10.4705 | -90.7738
10.4708 | -90.7743
10.4711 | -90.7747
10.4715 | -90.7754
10.4718 | -90.776
10.4721 | -90.7761
10.4724 | -90.7756
10.4727 | -90.7759
10.4731 | -90.777
10.4734 | -90.7775
10.4737 | -90.7781
10.474 | -90.7786
10.4743 | -90.7791
10.4747 | -90.7796
10.475 | -90.7802
10.4753 | -90.7809
10.4756 | -90.7814
10.476 | -90.782
10.4763 | -90.7825
10.4766 | -90.7827
10.4769 | -90.7834
10.4772 | -90.7845
10.4776 | -90.7849
10.4779 | -90.7857
10.4782 | -90.7863
10.4785 | -90.7868
10.4788 | -90.7873
10.4792 | -90.7878
10.4795 | -90.7885
10.4798 | -90.7892
10.4801 | -90.7897
10.4804 | -90.7899
10.4808 | -90.7903
10.4811 | -90.7908
10.4814 | -90.7913
10.4817 | -90.7918
10.4821 | -90.7925
10.4824 | -90.7937
10.4827 | -90.7947
10.483 | -90.7957
10.4833 | -90.7965
10.4837 | -90.7969
10.484 | -90.7977
10.4843 | -90.7983
10.4846 | -90.7987
10.4849 | -90.7991
10.4853 | -90.7996
10.4856 | -90.8005
10.4859 | -90.801
10.4862 | -90.8016
10.4866 | -90.802
10.4869 | -90.8023
10.4872 | -90.8029
10.4875 | -90.8035
10.4878 | -90.8047
10.4882 | -90.8059
10.4885 | -90.8063
10.4888 | -90.8065
10.4891 | -90.8068
10.4894 | -90.8072
10.4898 | -90.8073
10.4901 | -90.8076
10.4904 | -90.8085
10.4907 | -90.8089
10.4911 | -90.8092
10.4914 | -90.8097
10.4917 | -90.8102
10.492 | -90.8106
10.4923 | -90.8111
10.4927 | -90.8115
10.493 | -90.812
10.4933 | -90.8125
10.4936 | -90.8132
10.4939 | -90.8138
10.4943 | -90.8146
10.4946 | -90.8154
10.4949 | -90.8155
10.4952 | -90.8164
10.4956 | -90.817
10.4959 | -90.8175
10.4962 | -90.8182
10.4965 | -90.8187
10.4968 | -90.8191
10.4972 | -90.8197
10.4975 | -90.8201
10.4978 | -90.8205
10.4981 | -90.821
10.4984 | -90.8217
10.4988 | -90.8222
10.4991 | -90.8224
10.4994 | -90.8227
10.4997 | -90.8232
10.5001 | -90.8236
10.5004 | -90.8242
10.5007 | -90.8246
10.501 | -90.825
10.5013 | -90.8253
10.5017 | -90.8258
10.502 | -90.8265
10.5023 | -90.8276
10.5026 | -90.8284
10.5029 | -90.8288
10.5033 | -90.8295
10.5036 | -90.83
10.5039 | -90.8309
10.5042 | -90.8317
10.5046 | -90.8324
10.5049 | -90.833
10.5052 | -90.8331
10.5055 | -90.8333
10.5058 | -90.834
10.5062 | -90.8348
10.5065 | -90.8352
10.5068 | -90.8356
10.5071 | -90.836
10.5074 | -90.8365
10.5078 | -90.8371
10.5081 | -90.8376
10.5084 | -90.8383
10.5087 | -90.8395
10.509 | -90.8399
10.5094 | -90.8405
10.5097 | -90.841
10.51 | -90.8416
10.5103 | -90.8421
10.5107 | -90.8427
10.511 | -90.8435
10.5113 | -90.8441
10.5116 | -90.8442
10.5119 | -90.8448
10.5123 | -90.8456
10.5126 | -90.8462
10.5129 | -90.8465
10.5132 | -90.8466
10.5135 | -90.8468
10.5139 | -90.8472
10.5142 | -90.848
10.5145 | -90.8488
10.5148 | -90.8493
10.5152 | -90.8497
10.5155 | -90.8501
10.5158 | -90.8507
10.5161 | -90.8513
10.5164 | -90.8519
10.5168 | -90.8527
10.5171 | -90.8534
10.5174 | -90.854
10.5177 | -90.8544
10.518 | -90.855
10.5184 | -90.8557
10.5187 | -90.8557
10.519 | -90.8563
10.5193 | -90.8575
10.5197 | -90.8585
10.52 | -90.8595
10.5203 | -90.8603
10.5206 | -90.8607
10.5209 | -90.8611
10.5213 | -90.8615
10.5216 | -90.8621
10.5219 | -90.8627
10.5222 | -90.8636
10.5225 | -90.8643
10.5229 | -90.8649
10.5232 | -90.8653
10.5235 | -90.8658
10.5238 | -90.8661
10.5242 | -90.8665
10.5245 | -90.8664
10.5248 | -90.8665
10.5251 | -90.8669
10.5254 | -90.8679
10.5258 | -90.8697
10.5261 | -90.8701
10.5264 | -90.8704
10.5267 | -90.8708
10.527 | -90.8714
10.5274 | -90.8719
10.5277 | -90.8717
10.528 | -90.8732
10.5283 | -90.8741
10.5287 | -90.8747
10.529 | -90.875
10.5293 | -90.875
10.5296 | -90.8755
10.5299 | -90.8757
10.5303 | -90.8761
10.5306 | -90.8769
10.5309 | -90.8774
10.5312 | -90.8775
10.5315 | -90.8775
10.5319 | -90.8782
10.5322 | -90.8786
10.5325 | -90.879
10.5328 | -90.8795
10.5331 | -90.8798
10.5335 | -90.8804
10.5338 | -90.881
10.5341 | -90.8813
10.5344 | -90.8816
10.5348 | -90.8822
10.5351 | -90.8828
10.5354 | -90.8833
10.5357 | -90.8838
10.536 | -90.8843
10.5364 | -90.8844
10.5367 | -90.8845
10.537 | -90.885
10.5373 | -90.8854
10.5376 | -90.886
10.538 | -90.8865
10.5383 | -90.8874
10.5386 | -90.8879
10.5389 | -90.8889
10.5393 | -90.8906
10.5396 | -90.8914
10.5399 | -90.892
10.5402 | -90.893
10.5405 | -90.8939
10.5409 | -90.8948
10.5412 | -90.8956
10.5415 | -90.896
10.5418 | -90.8964
10.5421 | -90.8972
10.5425 | -90.8977
10.5428 | -90.8976
10.5431 | -90.8978
10.5434 | -90.8984
10.5438 | -90.899
10.5441 | -90.8995
10.5444 | -90.9
10.5447 | -90.9005
10.545 | -90.9011
10.5454 | -90.9017
10.5457 | -90.903
10.546 | -90.9042
10.5463 | -90.9044
10.5466 | -90.9049
10.547 | -90.905
10.5473 | -90.9052
10.5476 | -90.9052
10.5479 | -90.9057
10.5483 | -90.9062
10.5486 | -90.9067
10.5489 | -90.9074
10.5492 | -90.9077
10.5495 | -90.9083
10.5499 | -90.9093
10.5502 | -90.9104
10.5505 | -90.9111
10.5508 | -90.9117
10.5511 | -90.912
10.5515 | -90.9128
10.5518 | -90.9139
10.5521 | -90.9148
10.5524 | -90.9156
10.5527 | -90.9161
10.5531 | -90.9166
10.5534 | -90.9172
10.5537 | -90.9178
10.554 | -90.9185
10.5544 | -90.9186
10.5547 | -90.9183
10.555 | -90.9188
10.5553 | -90.9195
10.5556 | -90.92
10.556 | -90.9205
10.5563 | -90.921
10.5566 | -90.9213
10.5569 | -90.9217
10.5572 | -90.9227
10.5576 | -90.9238
10.5579 | -90.9243
10.5582 | -90.9251
10.5585 | -90.9257
10.5589 | -90.9267
10.5592 | -90.9275
10.5595 | -90.928
10.5598 | -90.9283
10.5601 | -90.929
10.5605 | -90.9298
10.5608 | -90.9304
10.5611 | -90.9309
10.5614 | -90.9316
10.5617 | -90.9319
10.5621 | -90.9322
10.5624 | -90.9326
10.5627 | -90.9331
10.563 | -90.9337
10.5634 | -90.934
10.5637 | -90.9343
10.564 | -90.9348
10.5643 | -90.9354
10.5646 | -90.9364
10.565 | -90.9376
10.5653 | -90.9382
10.5656 | -90.9387
10.5659 | -90.9391
10.5662 | -90.9398
10.5666 | -90.9402
10.5669 | -90.9406
10.5672 | -90.9413
10.5675 | -90.9417
10.5679 | -90.9421
10.5682 | -90.9429
10.5685 | -90.9436
10.5688 | -90.9441
10.5691 | -90.9448
10.5695 | -90.9456
10.5698 | -90.9456
10.5701 | -90.9459
10.5704 | -90.9467
10.5707 | -90.9472
10.5711 | -90.9476
10.5714 | -90.9481
10.5717 | -90.949
10.572 | -90.9494
10.5724 | -90.95
10.5727 | -90.9508
10.573 | -90.9512
10.5733 | -90.9519
10.5736 | -90.9528
10.574 | -90.9533
10.5743 | -90.9537
10.5746 | -90.9541
10.5749 | -90.9546
10.5752 | -90.9552
10.5756 | -90.9557
10.5759 | -90.9562
10.5762 | -90.9566
10.5765 | -90.9572
10.5768 | -90.9581
10.5772 | -90.9589
10.5775 | -90.9593
10.5778 | -90.9599
10.5781 | -90.9603
10.5785 | -90.9611
10.5788 | -90.9617
10.5791 | -90.9622
10.5794 | -90.963
10.5797 | -90.9642
10.5801 | -90.9651
10.5804 | -90.9659
10.5807 | -90.9669
10.581 | -90.9675
10.5813 | -90.9679
10.5817 | -90.9683
10.582 | -90.9687
10.5823 | -90.9692
10.5826 | -90.9704
10.583 | -90.9708
10.5833 | -90.9709
10.5836 | -90.9715
10.5839 | -90.9722
10.5842 | -90.9728
10.5846 | -90.9733
10.5849 | -90.9735
10.5852 | -90.9738
10.5855 | -90.9738
10.5858 | -90.9742
10.5862 | -90.9747
10.5865 | -90.9754
10.5868 | -90.976
10.5871 | -90.9766
10.5875 | -90.9773
10.5878 | -90.9785
10.5881 | -90.9795
10.5884 | -90.9795
10.5887 | -90.9798
10.5891 | -90.9808
10.5894 | -90.9813
10.5897 | -90.9818
10.59 | -90.9822
10.5903 | -90.9825
10.5907 | -90.983
10.591 | -90.9834
10.5913 | -90.9841
10.5916 | -90.9846
10.592 | -90.9852
10.5923 | -90.9861
10.5926 | -90.9869
10.5929 | -90.9874
10.5932 | -90.9875
10.5936 | -90.9879
10.5939 | -90.9885
10.5942 | -90.9896
10.5945 | -90.991
10.5948 | -90.9915
10.5952 | -90.9914
10.5955 | -90.991
10.5958 | -90.9914
10.5961 | -90.9919
10.5965 | -90.9924
10.5968 | -90.9932
10.5971 | -90.9938
10.5974 | -90.9949
10.5977 | -90.9962
10.5981 | -90.9969
10.5984 | -90.9978
10.5987 | -90.9988
10.599 | -90.999
10.5993 | -90.9995
10.5997 | -91.0002
10.6 | -91.0011
10.6003 | -91.0017
10.6006 | -91.0021
10.6009 | -91.0021
10.6013 | -91.0022
10.6016 | -91.0025
10.6019 | -91.0029
10.6022 | -91.0035
10.6026 | -91.0043
10.6029 | -91.0049
10.6032 | -91.0054
10.6035 | -91.0059
10.6038 | -91.0066
10.6042 | -91.0073
10.6045 | -91.008
10.6048 | -91.0089
10.6051 | -91.0095
10.6054 | -91.01
10.6058 | -91.0105
10.6061 | -91.0113
10.6064 | -91.0123
10.6067 | -91.0127
10.6071 | -91.013
10.6074 | -91.0133
10.6077 | -91.0138
10.608 | -91.0144
10.6083 | -91.0154
10.6087 | -91.0162
10.609 | -91.0168
10.6093 | -91.0176
10.6096 | -91.0185
10.6099 | -91.0195
10.6103 | -91.0202
10.6106 | -91.0203
10.6109 | -91.0208
10.6112 | -91.0218
10.6116 | -91.0226
10.6119 | -91.0231
10.6122 | -91.024
10.6125 | -91.0249
10.6128 | -91.0256
10.6132 | -91.0268
10.6135 | -91.028
10.6138 | -91.0287
10.6141 | -91.0295
10.6144 | -91.03
10.6148 | -91.03
10.6151 | -91.0299
10.6154 | -91.0298
10.6157 | -91.0299
10.6161 | -91.03
10.6164 | -91.0301
10.6167 | -91.0302
10.617 | -91.0306
10.6173 | -91.031
10.6177 | -91.0313
10.618 | -91.0317
10.6183 | -91.0322
10.6186 | -91.0328
10.6189 | -91.0336
10.6193 | -91.0342
10.6196 | -91.0348
10.6199 | -91.0354
10.6202 | -91.0361
10.6205 | -91.0367
10.6209 | -91.0373
10.6212 | -91.0373
10.6215 | -91.0371
10.6218 | -91.0377
10.6222 | -91.0384
10.6225 | -91.0389
10.6228 | -91.0395
10.6231 | -91.0398
10.6234 | -91.0406
10.6238 | -91.0417
10.6241 | -91.0425
10.6244 | -91.0434
10.6247 | -91.0441
10.625 | -91.0455
10.6254 | -91.047
10.6257 | -91.0477
10.626 | -91.0482
10.6263 | -91.0486
10.6267 | -91.0492
10.627 | -91.0499
10.6273 | -91.0506
10.6276 | -91.0513
10.6279 | -91.052
10.6283 | -91.0526
10.6286 | -91.0529
10.6289 | -91.0531
10.6292 | -91.0537
10.6295 | -91.0544
10.6299 | -91.0553
10.6302 | -91.0558
10.6305 | -91.0564
10.6308 | -91.0571
10.6312 | -91.0577
10.6315 | -91.0581
10.6318 | -91.0583
10.6321 | -91.0595
10.6324 | -91.0606
10.6328 | -91.0611
10.6331 | -91.0617
10.6334 | -91.0622
10.6337 | -91.0632
10.634 | -91.0634
10.6344 | -91.0635
10.6347 | -91.0639
10.635 | -91.0649
10.6353 | -91.0654
10.6357 | -91.066
10.636 | -91.0662
10.6363 | -91.0666
10.6366 | -91.0674
10.6369 | -91.068
10.6373 | -91.0684
10.6376 | -91.0691
10.6379 | -91.0696
10.6382 | -91.0702
10.6385 | -91.0708
10.6389 | -91.0714
10.6392 | -91.0719
10.6395 | -91.0724
10.6398 | -91.0726
10.6402 | -91.0732
10.6405 | -91.0737
10.6408 | -91.0743
10.6411 | -91.0752
10.6414 | -91.076
10.6418 | -91.077
10.6421 | -91.0776
10.6424 | -91.0779
10.6427 | -91.0785
10.643 | -91.0793
10.6434 | -91.08
10.6437 | -91.0804
10.644 | -91.0809
10.6443 | -91.0814
10.6446 | -91.0816
10.645 | -91.0819
10.6453 | -91.0824
10.6456 | -91.0831
10.6459 | -91.0838
10.6463 | -91.0844
10.6466 | -91.0849
10.6469 | -91.0855
10.6472 | -91.0861
10.6475 | -91.0867
10.6479 | -91.087
10.6482 | -91.0878
10.6485 | -91.0883
10.6488 | -91.0888
10.6491 | -91.0895
10.6495 | -91.0904
10.6498 | -91.0914
10.6501 | -91.092
10.6504 | -91.0924
10.6508 | -91.093
10.6511 | -91.0938
10.6514 | -91.0946
10.6517 | -91.0951
10.652 | -91.096
10.6524 | -91.0968
10.6527 | -91.098
10.653 | -91.0986
10.6533 | -91.0994
10.6536 | -91.1001
10.654 | -91.1006
10.6543 | -91.1014
10.6546 | -91.1023
10.6549 | -91.103
10.6553 | -91.1038
10.6556 | -91.1049
10.6559 | -91.1058
10.6562 | -91.1053
10.6565 | -91.1046
10.6569 | -91.1047
10.6572 | -91.1055
10.6575 | -91.1064
10.6578 | -91.1076
10.6581 | -91.1087
10.6585 | -91.1094
10.6588 | -91.1099
10.6591 | -91.1102
10.6594 | -91.1109
10.6598 | -91.1117
10.6601 | -91.1124
10.6604 | -91.1132
10.6607 | -91.1136
10.661 | -91.1141
10.6614 | -91.115
10.6617 | -91.1152
10.662 | -91.1159
10.6623 | -91.1166
10.6626 | -91.1174
10.663 | -91.1178
10.6633 | -91.1188
10.6636 | -91.1193
10.6639 | -91.1193
10.6642 | -91.1199
10.6646 | -91.1202
10.6649 | -91.121
10.6652 | -91.1223
10.6655 | -91.1236
10.6659 | -91.1241
10.6662 | -91.1246
10.6665 | -91.1254
10.6668 | -91.1256
10.6671 | -91.1257
10.6675 | -91.1263
10.6678 | -91.127
10.6681 | -91.1281
10.6684 | -91.1289
10.6687 | -91.1299
10.6691 | -91.1317
10.6694 | -91.1332
10.6697 | -91.1337
10.67 | -91.1344
10.6704 | -91.135
10.6707 | -91.1358
10.671 | -91.1368
10.6713 | -91.1372
10.6716 | -91.1377
10.672 | -91.1386
10.6723 | -91.1394
10.6726 | -91.1401
10.6729 | -91.1406
10.6732 | -91.1412
10.6736 | -91.1417
10.6739 | -91.1423
10.6742 | -91.143
10.6745 | -91.1439
10.6749 | -91.1443
10.6752 | -91.1449
10.6755 | -91.1454
10.6758 | -91.1461
10.6761 | -91.1467
10.6765 | -91.1473
10.6768 | -91.1482
10.6771 | -91.1489
10.6774 | -91.1495
10.6777 | -91.1502
10.6781 | -91.1508
10.6784 | -91.1514
10.6787 | -91.1523
10.679 | -91.1533
10.6794 | -91.1541
10.6797 | -91.1549
10.68 | -91.1556
10.6803 | -91.1565
10.6806 | -91.1571
10.681 | -91.1578
10.6813 | -91.1585
10.6816 | -91.1599
10.6819 | -91.1606
10.6822 | -91.1615
10.6826 | -91.163
10.6829 | -91.1637
10.6832 | -91.1644
10.6835 | -91.1649
10.6839 | -91.1657
10.6842 | -91.1664
10.6845 | -91.1669
10.6848 | -91.1674
10.6851 | -91.1679
10.6855 | -91.1693
10.6858 | -91.171
10.6861 | -91.1713
10.6864 | -91.1717
10.6867 | -91.1724
10.6871 | -91.1735
10.6874 | -91.1748
10.6877 | -91.1754
10.688 | -91.1762
10.6883 | -91.1779
10.6887 | -91.1795
10.689 | -91.18
10.6893 | -91.1807
10.6896 | -91.1829
10.69 | -91.1835
10.6903 | -91.1843
10.6906 | -91.1855
10.6909 | -91.1863
10.6912 | -91.1867
10.6916 | -91.1886
10.6919 | -91.1904
10.6922 | -91.1911
10.6925 | -91.1918
10.6928 | -91.1927
10.6932 | -91.1937
10.6935 | -91.1945
10.6938 | -91.1949
10.6941 | -91.1958
10.6945 | -91.1967
10.6948 | -91.1974
10.6951 | -91.198
10.6954 | -91.1985
10.6957 | -91.199
10.6961 | -91.1997
10.6964 | -91.2001
10.6967 | -91.2006
10.697 | -91.2011
10.6973 | -91.2016
10.6977 | -91.202
10.698 | -91.2024
10.6983 | -91.2036
10.6986 | -91.2045
10.699 | -91.2048
10.6993 | -91.2053
10.6996 | -91.2057
10.6999 | -91.2062
10.7002 | -91.2069
10.7006 | -91.207
10.7009 | -91.2075
10.7012 | -91.2079
10.7015 | -91.2084
10.7018 | -91.2089
10.7022 | -91.2092
10.7025 | -91.2095
10.7028 | -91.2101
10.7031 | -91.2106
10.7035 | -91.2112
10.7038 | -91.2122
10.7041 | -91.2128
10.7044 | -91.2136
10.7047 | -91.214
10.7051 | -91.2143
10.7054 | -91.2148
10.7057 | -91.2161
10.706 | -91.2172
10.7063 | -91.2176
10.7067 | -91.2177
10.707 | -91.2182
10.7073 | -91.2186
10.7076 | -91.219
10.708 | -91.22
10.7083 | -91.2212
10.7086 | -91.2223
10.7089 | -91.2225
10.7092 | -91.2228
10.7096 | -91.2237
10.7099 | -91.2243
10.7102 | -91.2248
10.7105 | -91.2254
10.7108 | -91.2265
10.7112 | -91.2274
10.7115 | -91.2281
10.7118 | -91.229
10.7121 | -91.2297
10.7125 | -91.2301
10.7128 | -91.2301
10.7131 | -91.2306
10.7134 | -91.2313
10.7137 | -91.2318
10.7141 | -91.232
10.7144 | -91.2332
10.7147 | -91.2341
10.715 | -91.235
10.7153 | -91.2357
10.7157 | -91.236
10.716 | -91.2367
10.7163 | -91.2374
10.7166 | -91.238
10.7169 | -91.2386
10.7173 | -91.2394
10.7176 | -91.2404
10.7179 | -91.2416
10.7182 | -91.2419
10.7186 | -91.2423
10.7189 | -91.2428
10.7192 | -91.2433
10.7195 | -91.244
10.7198 | -91.2447
10.7202 | -91.2456
10.7205 | -91.2467
10.7208 | -91.2472
10.7211 | -91.2476
10.7214 | -91.2483
10.7218 | -91.2492
10.7221 | -91.2505
10.7224 | -91.2522
10.7227 | -91.2533
10.7231 | -91.2537
10.7234 | -91.2542
10.7237 | -91.2548
10.724 | -91.2558
10.7243 | -91.257
10.7247 | -91.2579
10.725 | -91.2589
10.7253 | -91.2593
10.7256 | -91.2594
10.7259 | -91.26
10.7263 | -91.2607
10.7266 | -91.2614
10.7269 | -91.2617
10.7272 | -91.2624
10.7276 | -91.2634
10.7279 | -91.2643
10.7282 | -91.2649
10.7285 | -91.2653
10.7288 | -91.2655
10.7292 | -91.2661
10.7295 | -91.2666
10.7298 | -91.2676
10.7301 | -91.2683
10.7304 | -91.2686
10.7308 | -91.2689
10.7311 | -91.2694
10.7314 | -91.2697
10.7317 | -91.2703
10.7321 | -91.2707
10.7324 | -91.2703
10.7327 | -91.2706
10.733 | -91.2712
10.7333 | -91.2717
10.7337 | -91.2725
10.734 | -91.2732
10.7343 | -91.2737
10.7346 | -91.2738
10.7349 | -91.2743
10.7353 | -91.2747
10.7356 | -91.2752
10.7359 | -91.276
10.7362 | -91.2766
10.7365 | -91.2771
10.7369 | -91.2778
10.7372 | -91.2785
10.7375 | -91.2793
10.7378 | -91.2799
10.7382 | -91.2811
10.7385 | -91.2818
10.7388 | -91.2824
10.7391 | -91.283
10.7394 | -91.2835
10.7398 | -91.2841
10.7401 | -91.2847
10.7404 | -91.2852
10.7407 | -91.2859
10.741 | -91.2865
10.7414 | -91.2872
10.7417 | -91.2879
10.742 | -91.2887
10.7423 | -91.2895
10.7427 | -91.2899
10.743 | -91.2905
10.7433 | -91.2912
10.7436 | -91.292
10.7439 | -91.2927
10.7443 | -91.2934
10.7446 | -91.2941
10.7449 | -91.2951
10.7452 | -91.2957
10.7455 | -91.2961
10.7459 | -91.2967
10.7462 | -91.2973
10.7465 | -91.2976
10.7468 | -91.2983
10.7472 | -91.2991
10.7475 | -91.2995
10.7478 | -91.2999
10.7481 | -91.3006
10.7484 | -91.3013
10.7488 | -91.3018
10.7491 | -91.3021
10.7494 | -91.3026
10.7497 | -91.3031
10.75 | -91.3032
10.7504 | -91.304
10.7507 | -91.3049
10.751 | -91.3058
10.7513 | -91.3067
10.7517 | -91.3073
10.752 | -91.3079
10.7523 | -91.3086
10.7526 | -91.3092
10.7529 | -91.3098
10.7533 | -91.3099
10.7536 | -91.3104
10.7539 | -91.311
10.7542 | -91.3107
10.7545 | -91.3115
10.7549 | -91.3122
10.7552 | -91.3128
10.7555 | -91.3138
10.7558 | -91.3148
10.7561 | -91.3156
10.7565 | -91.3163
10.7568 | -91.317
10.7571 | -91.3178
10.7574 | -91.3188
10.7578 | -91.3199
10.7581 | -91.3208
10.7584 | -91.3217
10.7587 | -91.3227
10.759 | -91.3232
10.7594 | -91.3236
10.7597 | -91.3242
10.76 | -91.325
10.7603 | -91.3254
10.7606 | -91.3263
10.761 | -91.3272
10.7613 | -91.328
10.7616 | -91.3287
10.7619 | -91.3293
10.7623 | -91.33
10.7626 | -91.33
10.7629 | -91.3304
10.7632 | -91.3313
10.7635 | -91.3317
10.7639 | -91.3324
10.7642 | -91.3331
10.7645 | -91.3338
10.7648 | -91.3346
10.7651 | -91.3349
10.7655 | -91.3355
10.7658 | -91.3366
10.7661 | -91.3371
10.7664 | -91.337
10.7668 | -91.3372
10.7671 | -91.3379
10.7674 | -91.3388
10.7677 | -91.3396
10.768 | -91.3401
10.7684 | -91.3405
10.7687 | -91.3414
10.769 | -91.3429
10.7693 | -91.3439
10.7696 | -91.3445
10.77 | -91.3454
10.7703 | -91.3463
10.7706 | -91.3471
10.7709 | -91.3477
10.7713 | -91.3482
10.7716 | -91.3488
10.7719 | -91.3496
10.7722 | -91.3506
10.7725 | -91.3514
10.7729 | -91.3523
10.7732 | -91.353
10.7735 | -91.354
10.7738 | -91.355
10.7741 | -91.356
10.7745 | -91.3567
10.7748 | -91.3575
10.7751 | -91.3577
10.7754 | -91.3586
10.7758 | -91.3593
10.7761 | -91.3604
10.7764 | -91.362
10.7767 | -91.3628
10.777 | -91.3638
10.7774 | -91.365
10.7777 | -91.3655
10.778 | -91.3666
10.7783 | -91.3674
10.7786 | -91.3678
10.779 | -91.3684
10.7793 | -91.3695
10.7796 | -91.3714
10.7799 | -91.373
10.7803 | -91.3737
10.7806 | -91.3742
10.7809 | -91.3745
10.7812 | -91.3749
10.7815 | -91.3756
10.7819 | -91.3762
10.7822 | -91.3769
10.7825 | -91.3775
10.7828 | -91.3784
10.7831 | -91.3795
10.7835 | -91.3807
10.7838 | -91.3815
10.7841 | -91.3829
10.7844 | -91.384
10.7847 | -91.3851
10.7851 | -91.386
10.7854 | -91.3867
10.7857 | -91.3875
10.786 | -91.3884
10.7864 | -91.3892
10.7867 | -91.3894
10.787 | -91.3899
10.7873 | -91.3907
10.7876 | -91.3914
10.788 | -91.3922
10.7883 | -91.3928
10.7886 | -91.3935
10.7889 | -91.3941
10.7892 | -91.3947
10.7896 | -91.3955
10.7899 | -91.3961
10.7902 | -91.3968
10.7905 | -91.397
10.7909 | -91.3971
10.7912 | -91.3979
10.7915 | -91.3985
10.7918 | -91.399
10.7921 | -91.3997
10.7925 | -91.4004
10.7928 | -91.4011
10.7931 | -91.4018
10.7934 | -91.4026
10.7937 | -91.4035
10.7941 | -91.4041
10.7944 | -91.4047
10.7947 | -91.4052
10.795 | -91.406
10.7954 | -91.4068
10.7957 | -91.4074
10.796 | -91.4081
10.7963 | -91.4089
10.7966 | -91.4096
10.797 | -91.4104
10.7973 | -91.4112
10.7976 | -91.4123
10.7979 | -91.4129
10.7982 | -91.4134
10.7986 | -91.4138
10.7989 | -91.4145
10.7992 | -91.4148
10.7995 | -91.4153
10.7999 | -91.4159
10.8002 | -91.4165
10.8005 | -91.4172
10.8008 | -91.418
10.8011 | -91.4192
10.8015 | -91.4204
10.8018 | -91.4214
10.8021 | -91.4226
10.8024 | -91.4231
10.8027 | -91.4236
10.8031 | -91.4245
10.8034 | -91.4252
10.8037 | -91.4258
10.804 | -91.4265
10.8043 | -91.4271
10.8047 | -91.4283
10.805 | -91.4295
10.8053 | -91.4304
10.8056 | -91.431
10.806 | -91.4314
10.8063 | -91.4319
10.8066 | -91.4329
10.8069 | -91.4339
10.8072 | -91.4357
10.8076 | -91.4371
10.8079 | -91.4387
10.8082 | -91.4398
10.8085 | -91.4398
10.8088 | -91.4398
10.8092 | -91.4404
10.8095 | -91.4411
10.8098 | -91.4415
10.8101 | -91.4419
10.8105 | -91.4423
10.8108 | -91.4429
10.8111 | -91.4436
10.8114 | -91.4443
10.8117 | -91.4451
10.8121 | -91.4457
10.8124 | -91.4463
10.8127 | -91.4471
10.813 | -91.4478
10.8133 | -91.4487
10.8137 | -91.4497
10.814 | -91.4502
10.8143 | -91.4517
10.8146 | -91.4525
10.815 | -91.453
10.8153 | -91.4537
10.8156 | -91.4554
10.8159 | -91.4564
10.8162 | -91.4572
10.8166 | -91.4579
10.8169 | -91.4582
10.8172 | -91.4585
10.8175 | -91.4586
10.8178 | -91.459
10.8182 | -91.4597
10.8185 | -91.4609
10.8188 | -91.4619
10.8191 | -91.4625
10.8195 | -91.4627
10.8198 | -91.4628
10.8201 | -91.4633
10.8204 | -91.4638
10.8207 | -91.4645
10.8211 | -91.4654
10.8214 | -91.4659
10.8217 | -91.4668
10.822 | -91.4676
10.8223 | -91.4684
10.8227 | -91.4697
10.823 | -91.4705
10.8233 | -91.4715
10.8236 | -91.4724
10.824 | -91.4732
10.8243 | -91.4744
10.8246 | -91.4748
10.8249 | -91.4751
10.8252 | -91.4759
10.8256 | -91.4766
10.8259 | -91.4768
10.8262 | -91.4773
10.8265 | -91.4777
10.8268 | -91.4782
10.8272 | -91.4785
10.8275 | -91.4791
10.8278 | -91.4796
10.8281 | -91.4802
10.8284 | -91.4801
10.8288 | -91.4804
10.8291 | -91.4808
10.8294 | -91.4816
10.8297 | -91.4824
10.8301 | -91.4831
10.8304 | -91.4843
10.8307 | -91.485
10.831 | -91.4854
10.8313 | -91.4859
10.8317 | -91.4868
10.832 | -91.4876
10.8323 | -91.4886
10.8326 | -91.4895
10.8329 | -91.4907
10.8333 | -91.4917
10.8336 | -91.4925
10.8339 | -91.4931
10.8342 | -91.4938
10.8346 | -91.4945
10.8349 | -91.4952
10.8352 | -91.4957
10.8355 | -91.4966
10.8358 | -91.4976
10.8362 | -91.4986
10.8365 | -91.4988
10.8368 | -91.4989
10.8371 | -91.4996
10.8374 | -91.5002
10.8378 | -91.5008
10.8381 | -91.5013
10.8384 | -91.5022
10.8387 | -91.5029
10.8391 | -91.5035
10.8394 | -91.5042
10.8397 | -91.5049
10.84 | -91.5055
10.8403 | -91.5063
10.8407 | -91.5074
10.841 | -91.5082
10.8413 | -91.5088
10.8416 | -91.509
10.8419 | -91.5097
10.8423 | -91.5106
10.8426 | -91.5117
10.8429 | -91.5131
10.8432 | -91.5136
10.8436 | -91.5148
10.8439 | -91.5154
10.8442 | -91.516
10.8445 | -91.5165
10.8448 | -91.5173
10.8452 | -91.5179
10.8455 | -91.5186
10.8458 | -91.5195
10.8461 | -91.5199
10.8464 | -91.5203
10.8468 | -91.5209
10.8471 | -91.5218
10.8474 | -91.523
10.8477 | -91.5238
10.8481 | -91.5245
10.8484 | -91.525
10.8487 | -91.5262
10.849 | -91.5269
10.8493 | -91.5276
10.8497 | -91.5284
10.85 | -91.5292
10.8503 | -91.53
10.8506 | -91.5306
10.8509 | -91.5311
10.8513 | -91.5315
10.8516 | -91.5321
10.8519 | -91.5329
10.8522 | -91.5339
10.8525 | -91.5355
10.8529 | -91.5375
10.8532 | -91.5382
10.8535 | -91.5389
10.8538 | -91.5399
10.8542 | -91.5406
10.8545 | -91.5418
10.8548 | -91.5423
10.8551 | -91.5424
10.8554 | -91.5433
10.8558 | -91.5439
10.8561 | -91.5446
10.8564 | -91.5455
10.8567 | -91.5462
10.857 | -91.5463
10.8574 | -91.546
10.8577 | -91.5461
10.858 | -91.5472
10.8583 | -91.548
10.8587 | -91.5488
10.859 | -91.5499
10.8593 | -91.5512
10.8596 | -91.5521
10.8599 | -91.5529
10.8603 | -91.5535
10.8606 | -91.5535
10.8609 | -91.5543
10.8612 | -91.5551
10.8615 | -91.5558
10.8619 | -91.5564
10.8622 | -91.5565
10.8625 | -91.5566
10.8628 | -91.5582
10.8632 | -91.5597
10.8635 | -91.5607
10.8638 | -91.5617
10.8641 | -91.5618
10.8644 | -91.5622
10.8648 | -91.5629
10.8651 | -91.5634
10.8654 | -91.5639
10.8657 | -91.5648
10.866 | -91.5655
10.8664 | -91.5659
10.8667 | -91.5666
10.867 | -91.5672
10.8673 | -91.5677
10.8677 | -91.5679
10.868 | -91.5687
10.8683 | -91.5693
10.8686 | -91.5693
10.8689 | -91.5697
10.8693 | -91.5709
10.8696 | -91.5721
10.8699 | -91.5727
10.8702 | -91.5731
10.8705 | -91.5734
10.8709 | -91.574
10.8712 | -91.5748
10.8715 | -91.5748
10.8718 | -91.5748
10.8721 | -91.5732
10.8725 | -91.5729
10.8728 | -91.5698
10.8731 | -91.5663
10.8734 | -91.5649
10.8738 | -91.5591
10.8741 | -91.5524
10.8744 | -91.5508
10.8747 | -91.5505
10.875 | -91.5491
10.8754 | -91.5488
10.8757 | -91.5487
10.876 | -91.5485
10.8763 | -91.5481
10.8766 | -91.5487
10.877 | -91.5505
10.8773 | -91.5518
10.8776 | -91.5527
10.8779 | -91.5534
10.8783 | -91.5539
10.8786 | -91.5545
10.8789 | -91.5552
10.8792 | -91.5557
10.8795 | -91.5561
10.8799 | -91.5564
10.8802 | -91.5561
10.8805 | -91.5565
10.8808 | -91.557
10.8811 | -91.5575
10.8815 | -91.5582
10.8818 | -91.5587
10.8821 | -91.5589
10.8824 | -91.5589
10.8828 | -91.5591
10.8831 | -91.5598
10.8834 | -91.5603
10.8837 | -91.5609
10.884 | -91.561
10.8844 | -91.5613
10.8847 | -91.5617
10.885 | -91.5617
10.8853 | -91.5616
10.8856 | -91.5628
10.886 | -91.5639
10.8863 | -91.5641
10.8866 | -91.5637
10.8869 | -91.5639
10.8873 | -91.5643
10.8876 | -91.5647
10.8879 | -91.5653
10.8882 | -91.5657
10.8885 | -91.5664
10.8889 | -91.5672
10.8892 | -91.5677
10.8895 | -91.5679
10.8898 | -91.5675
10.8901 | -91.5682
10.8905 | -91.5689
10.8908 | -91.5697
10.8911 | -91.5699
10.8914 | -91.5702
10.8918 | -91.5705
10.8921 | -91.571
10.8924 | -91.572
10.8927 | -91.5737
10.893 | -91.5746
10.8934 | -91.575
10.8937 | -91.5753
10.894 | -91.5757
10.8943 | -91.5762
10.8946 | -91.5767
10.895 | -91.5781
10.8953 | -91.5798
10.8956 | -91.5803
10.8959 | -91.5809
10.8962 | -91.5816
10.8966 | -91.5827
10.8969 | -91.5835
10.8972 | -91.5838
10.8975 | -91.584
10.8979 | -91.5844
10.8982 | -91.5847
10.8985 | -91.585
10.8988 | -91.5854
10.8991 | -91.5858
10.8995 | -91.5859
10.8998 | -91.5862
10.9001 | -91.5867
10.9004 | -91.5867
10.9007 | -91.5862
10.9011 | -91.5865
10.9014 | -91.587
10.9017 | -91.5877
10.902 | -91.588
10.9024 | -91.5883
10.9027 | -91.5887
10.903 | -91.5889
10.9033 | -91.5889
10.9036 | -91.5894
10.904 | -91.5896
10.9043 | -91.59
10.9046 | -91.5903
10.9049 | -91.5905
10.9052 | -91.5911
10.9056 | -91.5918
10.9059 | -91.5922
10.9062 | -91.5928
10.9065 | -91.593
10.9069 | -91.5933
10.9072 | -91.5937
10.9075 | -91.5941
10.9078 | -91.5946
10.9081 | -91.5951
10.9085 | -91.5958
10.9088 | -91.5966
10.9091 | -91.5972
10.9094 | -91.5976
10.9097 | -91.598
10.9101 | -91.5983
10.9104 | -91.5987
10.9107 | -91.5992
10.911 | -91.5994
10.9114 | -91.6
10.9117 | -91.6003
10.912 | -91.6006
10.9123 | -91.6013
10.9126 | -91.602
10.913 | -91.6022
10.9133 | -91.6024
10.9136 | -91.6026
10.9139 | -91.6033
10.9142 | -91.604
10.9146 | -91.6047
10.9149 | -91.6046
10.9152 | -91.6046
10.9155 | -91.6048
10.9159 | -91.6052
10.9162 | -91.6056
10.9165 | -91.6061
10.9168 | -91.6063
10.9171 | -91.6068
10.9175 | -91.6074
10.9178 | -91.6073
10.9181 | -91.6082
10.9184 | -91.6088
10.9187 | -91.6093
10.9191 | -91.6097
10.9194 | -91.6097
10.9197 | -91.6099
10.92 | -91.6097
10.9203 | -91.6098
10.9207 | -91.6101
10.921 | -91.6103
10.9213 | -91.6106
10.9216 | -91.6113
10.922 | -91.6122
10.9223 | -91.6128
10.9226 | -91.6131
10.9229 | -91.6133
10.9232 | -91.6135
10.9236 | -91.6135
10.9239 | -91.6137
10.9242 | -91.6145
10.9245 | -91.6154
10.9248 | -91.6159
10.9252 | -91.6161
10.9255 | -91.6164
10.9258 | -91.6164
10.9261 | -91.6165
10.9265 | -91.617
10.9268 | -91.6173
10.9271 | -91.6176
10.9274 | -91.6184
10.9277 | -91.6188
10.9281 | -91.6192
10.9284 | -91.6205
10.9287 | -91.6216
10.929 | -91.622
10.9293 | -91.6222
10.9297 | -91.6225
10.93 | -91.6227
10.9303 | -91.6231
10.9306 | -91.6235
10.931 | -91.6236
10.9313 | -91.6238
10.9316 | -91.624
10.9319 | -91.6244
10.9322 | -91.6245
10.9326 | -91.6244
10.9329 | -91.6254
10.9332 | -91.6264
10.9335 | -91.6266
10.9338 | -91.6269
10.9342 | -91.6269
10.9345 | -91.6269
10.9348 | -91.6275
10.9351 | -91.6281
10.9355 | -91.6284
10.9358 | -91.6281
10.9361 | -91.6282
10.9364 | -91.6285
10.9367 | -91.6288
10.9371 | -91.6291
10.9374 | -91.6293
10.9377 | -91.6289
10.938 | -91.6288
10.9383 | -91.6287
10.9387 | -91.6289
10.939 | -91.6295
10.9393 | -91.63
10.9396 | -91.6303
10.9399 | -91.6304
10.9403 | -91.6305
10.9406 | -91.6307
10.9409 | -91.631
10.9412 | -91.6315
10.9416 | -91.6317
10.9419 | -91.6319
10.9422 | -91.632
10.9425 | -91.632
10.9428 | -91.6321
10.9432 | -91.6317
10.9435 | -91.6315
10.9438 | -91.6317
10.9441 | -91.632
10.9444 | -91.6325
10.9448 | -91.6332
10.9451 | -91.6334
10.9454 | -91.6334
10.9457 | -91.6335
10.9461 | -91.6338
10.9464 | -91.6338
10.9467 | -91.6338
10.947 | -91.6334
10.9473 | -91.6333
10.9477 | -91.6334
10.948 | -91.6333
10.9483 | -91.6337
10.9486 | -91.6342
10.9489 | -91.6341
10.9493 | -91.6348
10.9496 | -91.6355
10.9499 | -91.6354
10.9502 | -91.6356
10.9506 | -91.6363
10.9509 | -91.6368
10.9512 | -91.6372
10.9515 | -91.6374
10.9518 | -91.6371
10.9522 | -91.6371
10.9525 | -91.6373
10.9528 | -91.6379
10.9531 | -91.6383
10.9534 | -91.6385
10.9538 | -91.6384
10.9541 | -91.6386
10.9544 | -91.6392
10.9547 | -91.6392
10.9551 | -91.6389
10.9554 | -91.6391
10.9557 | -91.6394
10.956 | -91.6392
10.9563 | -91.6388
10.9567 | -91.6388
10.957 | -91.6387
10.9573 | -91.6393
10.9576 | -91.6395
10.9579 | -91.6397
10.9583 | -91.6398
10.9586 | -91.6396
10.9589 | -91.6393
10.9592 | -91.6391
10.9596 | -91.6394
10.9599 | -91.6397
10.9602 | -91.6395
10.9605 | -91.64
10.9608 | -91.6403
10.9612 | -91.6405
10.9615 | -91.6406
10.9618 | -91.6406
10.9621 | -91.6408
10.9624 | -91.6408
10.9628 | -91.6408
10.9631 | -91.6407
10.9634 | -91.6405
10.9637 | -91.6403
10.9641 | -91.6401
10.9644 | -91.6401
10.9647 | -91.6403
10.965 | -91.6407
10.9653 | -91.641
10.9657 | -91.6411
10.966 | -91.6413
10.9663 | -91.6419
10.9666 | -91.6417
10.9669 | -91.6417
10.9673 | -91.6422
10.9676 | -91.6421
10.9679 | -91.6416
10.9682 | -91.6413
10.9685 | -91.641
10.9689 | -91.6409
10.9692 | -91.6406
10.9695 | -91.6401
10.9698 | -91.6396
10.9702 | -91.6393
10.9705 | -91.6391
10.9708 | -91.6391
10.9711 | -91.6393
10.9714 | -91.6392
10.9718 | -91.6389
10.9721 | -91.639
10.9724 | -91.6391
10.9727 | -91.6389
10.973 | -91.6382
10.9734 | -91.6382
10.9737 | -91.6381
10.974 | -91.6379
10.9743 | -91.638
10.9747 | -91.6382
10.975 | -91.6382
10.9753 | -91.6379
10.9756 | -91.6377
10.9759 | -91.6367
10.9763 | -91.6361
10.9766 | -91.6359
10.9769 | -91.6354
10.9772 | -91.6352
10.9775 | -91.6349
10.9779 | -91.6351
10.9782 | -91.6349
10.9785 | -91.6351
10.9788 | -91.635
10.9792 | -91.6348
10.9795 | -91.6347
10.9798 | -91.6347
10.9801 | -91.6344
10.9804 | -91.6343
10.9808 | -91.6342
10.9811 | -91.634
10.9814 | -91.6341
10.9817 | -91.6343
10.982 | -91.6343
10.9824 | -91.6343
10.9827 | -91.6342
10.983 | -91.634
10.9833 | -91.6341
10.9837 | -91.6346
10.984 | -91.6349
10.9843 | -91.6348
10.9846 | -91.6352
10.9849 | -91.6355
10.9853 | -91.6356
10.9856 | -91.6353
10.9859 | -91.6353
10.9862 | -91.6345
10.9865 | -91.6346
10.9869 | -91.6353
10.9872 | -91.6357
10.9875 | -91.6356
10.9878 | -91.6352
10.9881 | -91.6355
10.9885 | -91.6353
10.9888 | -91.6355
10.9891 | -91.6355
10.9894 | -91.6356
10.9898 | -91.6359
10.9901 | -91.636
10.9904 | -91.6356
10.9907 | -91.6351
10.991 | -91.6352
10.9914 | -91.6352
10.9917 | -91.635
10.992 | -91.6351
10.9923 | -91.635
10.9926 | -91.635
10.993 | -91.635
10.9933 | -91.635
10.9936 | -91.6353
10.9939 | -91.6353
10.9943 | -91.6352
10.9946 | -91.6352
10.9949 | -91.6356
10.9952 | -91.6362
10.9955 | -91.6359
10.9959 | -91.6358
10.9962 | -91.636
10.9965 | -91.6361
10.9968 | -91.6361
10.9971 | -91.6355
10.9975 | -91.6354
10.9978 | -91.6358
10.9981 | -91.6368
10.9984 | -91.637
10.9988 | -91.6371
10.9991 | -91.6371
10.9994 | -91.6367
10.9997 | -91.6365
11 | -91.6365
11.0004 | -91.6364
11.0007 | -91.6364
11.001 | -91.6363
11.0013 | -91.6361
11.0016 | -91.6358
11.002 | -91.6359
11.0023 | -91.6363
11.0026 | -91.6362
11.0029 | -91.6359
11.0033 | -91.6358
11.0036 | -91.6359
11.0039 | -91.6358
11.0042 | -91.6355
11.0045 | -91.6344
11.0049 | -91.634
11.0052 | -91.6337
11.0055 | -91.6335
11.0058 | -91.6333
11.0061 | -91.6332
11.0065 | -91.6326
11.0068 | -91.6322
11.0071 | -91.632
11.0074 | -91.6321
11.0077 | -91.6319
11.0081 | -91.6317
11.0084 | -91.6314
11.0087 | -91.6308
11.009 | -91.6306
11.0094 | -91.6305
11.0097 | -91.6302
11.01 | -91.6301
11.0103 | -91.63
11.0106 | -91.6302
11.011 | -91.63
11.0113 | -91.6297
11.0116 | -91.6294
11.0119 | -91.6287
11.0122 | -91.6283
11.0126 | -91.6284
11.0129 | -91.6282
11.0132 | -91.6282
11.0135 | -91.6281
11.0139 | -91.6277
11.0142 | -91.6276
11.0145 | -91.6284
11.0148 | -91.6286
11.0151 | -91.6284
11.0155 | -91.6286
11.0158 | -91.629
11.0161 | -91.6294
11.0164 | -91.6293
11.0167 | -91.6292
11.0171 | -91.6289
11.0174 | -91.6293
11.0177 | -91.6298
11.018 | -91.6297
11.0184 | -91.6293
11.0187 | -91.629
11.019 | -91.6286
11.0193 | -91.6281
11.0196 | -91.6275
11.02 | -91.6271
11.0203 | -91.6263
11.0206 | -91.6258
11.0209 | -91.6255
11.0212 | -91.6253
11.0216 | -91.6249
11.0219 | -91.6244
11.0222 | -91.6241
11.0225 | -91.6239
11.0229 | -91.6234
11.0232 | -91.622
11.0235 | -91.6219
11.0238 | -91.6215
11.0241 | -91.6215
11.0245 | -91.6218
11.0248 | -91.6216
11.0251 | -91.6214
11.0254 | -91.6217
11.0257 | -91.6221
11.0261 | -91.6215
11.0264 | -91.6208
11.0267 | -91.6206
11.027 | -91.6205
11.0274 | -91.6202
11.0277 | -91.62
11.028 | -91.6199
11.0283 | -91.6197
11.0286 | -91.6193
11.029 | -91.6186
11.0293 | -91.6188
11.0296 | -91.619
11.0299 | -91.6192
11.0302 | -91.6197
11.0306 | -91.619
11.0309 | -91.6193
11.0312 | -91.6191
11.0315 | -91.6189
11.0319 | -91.6186
11.0322 | -91.6184
11.0325 | -91.618
11.0328 | -91.618
11.0331 | -91.6176
11.0335 | -91.6173
11.0338 | -91.617
11.0341 | -91.6171
11.0344 | -91.6168
11.0347 | -91.6163
11.0351 | -91.6159
11.0354 | -91.6156
11.0357 | -91.6155
11.036 | -91.6152
11.0363 | -91.6145
11.0367 | -91.6135
11.037 | -91.6128
11.0373 | -91.6125
11.0376 | -91.6118
11.038 | -91.6112
11.0383 | -91.6109
11.0386 | -91.6106
11.0389 | -91.61
11.0392 | -91.6093
11.0396 | -91.6086
11.0399 | -91.6086
11.0402 | -91.6086
11.0405 | -91.6079
11.0408 | -91.6073
11.0412 | -91.6071
11.0415 | -91.6067
11.0418 | -91.606
11.0421 | -91.6057
11.0425 | -91.6058
11.0428 | -91.6055
11.0431 | -91.6051
11.0434 | -91.6049
11.0437 | -91.6045
11.0441 | -91.604
11.0444 | -91.6036
11.0447 | -91.6035
11.045 | -91.6033
11.0453 | -91.6029
11.0457 | -91.6026
11.046 | -91.6026
11.0463 | -91.6026
11.0466 | -91.6021
11.047 | -91.6017
11.0473 | -91.6013
11.0476 | -91.601
11.0479 | -91.6006
11.0482 | -91.6002
11.0486 | -91.5997
11.0489 | -91.5995
11.0492 | -91.6004
11.0495 | -91.601
11.0498 | -91.6009
11.0502 | -91.6002
11.0505 | -91.5999
11.0508 | -91.5994
11.0511 | -91.5986
11.0515 | -91.5982
11.0518 | -91.598
11.0521 | -91.5976
11.0524 | -91.5971
11.0527 | -91.5961
11.0531 | -91.5956
11.0534 | -91.5951
11.0537 | -91.5947
11.054 | -91.594
11.0543 | -91.5934
11.0547 | -91.593
11.055 | -91.5928
11.0553 | -91.593
11.0556 | -91.5924
11.0559 | -91.5916
11.0563 | -91.5912
11.0566 | -91.5907
11.0569 | -91.5903
11.0572 | -91.59
11.0576 | -91.5897
11.0579 | -91.5896
11.0582 | -91.5895
11.0585 | -91.5894
11.0588 | -91.5892
11.0592 | -91.5888
11.0595 | -91.5884
11.0598 | -91.5885
11.0601 | -91.588
11.0604 | -91.5884
11.0608 | -91.5883
11.0611 | -91.5877
11.0614 | -91.5872
11.0617 | -91.5867
11.0621 | -91.5867
11.0624 | -91.5866
11.0627 | -91.5868
11.063 | -91.5867
11.0633 | -91.5865
11.0637 | -91.5861
11.064 | -91.5857
11.0643 | -91.5849
11.0646 | -91.5843
11.0649 | -91.584
11.0653 | -91.5834
11.0656 | -91.5829
11.0659 | -91.5831
11.0662 | -91.5828
11.0666 | -91.5826
11.0669 | -91.5821
11.0672 | -91.5815
11.0675 | -91.5804
11.0678 | -91.5799
11.0682 | -91.5795
11.0685 | -91.5789
11.0688 | -91.5783
11.0691 | -91.5775
11.0694 | -91.5766
11.0698 | -91.5761
11.0701 | -91.5757
11.0704 | -91.5753
11.0707 | -91.5749
11.0711 | -91.5743
11.0714 | -91.5737
11.0717 | -91.573
11.072 | -91.5728
11.0723 | -91.5726
11.0727 | -91.5723
11.073 | -91.5717
11.0733 | -91.5712
11.0736 | -91.5714
11.0739 | -91.5714
11.0743 | -91.571
11.0746 | -91.5705
11.0749 | -91.57
11.0752 | -91.5693
11.0756 | -91.5688
11.0759 | -91.5683
11.0762 | -91.5675
11.0765 | -91.5668
11.0768 | -91.5663
11.0772 | -91.5659
11.0775 | -91.5655
11.0778 | -91.5651
11.0781 | -91.5646
11.0784 | -91.564
11.0788 | -91.5636
11.0791 | -91.5631
11.0794 | -91.5626
11.0797 | -91.5619
11.08 | -91.561
11.0804 | -91.5602
11.0807 | -91.5597
11.081 | -91.5592
11.0813 | -91.5588
11.0817 | -91.5583
11.082 | -91.5583
11.0823 | -91.5578
11.0826 | -91.5573
11.0829 | -91.5569
11.0833 | -91.5561
11.0836 | -91.5559
11.0839 | -91.5558
11.0842 | -91.5552
11.0845 | -91.5548
11.0849 | -91.5545
11.0852 | -91.5539
11.0855 | -91.5531
11.0858 | -91.5524
11.0862 | -91.552
11.0865 | -91.552
11.0868 | -91.5518
11.0871 | -91.5517
11.0874 | -91.5512
11.0878 | -91.5507
11.0881 | -91.5497
11.0884 | -91.549
11.0887 | -91.5484
11.089 | -91.5477
11.0894 | -91.5475
11.0897 | -91.5471
11.09 | -91.5465
11.0903 | -91.5459
11.0907 | -91.5453
11.091 | -91.5449
11.0913 | -91.5444
11.0916 | -91.5439
11.0919 | -91.5434
11.0923 | -91.5429
11.0926 | -91.543
11.0929 | -91.5422
11.0932 | -91.5419
11.0935 | -91.5415
11.0939 | -91.541
11.0942 | -91.541
11.0945 | -91.5401
11.0948 | -91.5395
11.0952 | -91.539
11.0955 | -91.5386
11.0958 | -91.5382
11.0961 | -91.5376
11.0964 | -91.5373
11.0968 | -91.5369
11.0971 | -91.537
11.0974 | -91.5365
11.0977 | -91.5365
11.098 | -91.5357
11.0984 | -91.5347
11.0987 | -91.5341
11.099 | -91.5346
11.0993 | -91.535
11.0997 | -91.5345
11.1 | -91.5341
11.1003 | -91.5341
11.1006 | -91.5341
11.1009 | -91.5336
11.1013 | -91.5333
11.1016 | -91.5327
11.1019 | -91.5324
11.1022 | -91.5321
11.1025 | -91.5321
11.1029 | -91.5323
11.1032 | -91.5318
11.1035 | -91.5312
11.1038 | -91.5308
11.1041 | -91.5301
11.1045 | -91.5296
11.1048 | -91.5291
11.1051 | -91.5284
11.1054 | -91.528
11.1058 | -91.5277
11.1061 | -91.5275
11.1064 | -91.5274
11.1067 | -91.527
11.107 | -91.5261
11.1074 | -91.5256
11.1077 | -91.5248
11.108 | -91.524
11.1083 | -91.5231
11.1086 | -91.5224
11.109 | -91.5217
11.1093 | -91.5212
11.1096 | -91.5207
11.1099 | -91.5203
11.1103 | -91.5193
11.1106 | -91.5184
11.1109 | -91.5177
11.1112 | -91.5171
11.1115 | -91.5165
11.1119 | -91.5154
11.1122 | -91.5145
11.1125 | -91.5144
11.1128 | -91.5147
11.1131 | -91.5143
11.1135 | -91.5139
11.1138 | -91.5136
11.1141 | -91.513
11.1144 | -91.5124
11.1148 | -91.5121
11.1151 | -91.5117
11.1154 | -91.5115
11.1157 | -91.5109
11.116 | -91.5105
11.1164 | -91.51
11.1167 | -91.5096
11.117 | -91.5091
11.1173 | -91.5087
11.1176 | -91.5081
11.118 | -91.5079
11.1183 | -91.5079
11.1186 | -91.5076
11.1189 | -91.507
11.1193 | -91.5064
11.1196 | -91.5062
11.1199 | -91.5059
11.1202 | -91.5057
11.1205 | -91.5055
11.1209 | -91.5051
11.1212 | -91.505
11.1215 | -91.5049
11.1218 | -91.5043
11.1221 | -91.5036
11.1225 | -91.503
11.1228 | -91.5022
11.1231 | -91.5015
11.1234 | -91.501
11.1237 | -91.5004
11.1241 | -91.4998
11.1244 | -91.4996
11.1247 | -91.499
11.125 | -91.4986
11.1254 | -91.4979
11.1257 | -91.4974
11.126 | -91.4969
11.1263 | -91.496
11.1266 | -91.4952
11.127 | -91.4944
11.1273 | -91.4937
11.1276 | -91.4933
11.1279 | -91.4928
11.1282 | -91.4926
11.1286 | -91.4927
11.1289 | -91.4925
11.1292 | -91.4921
11.1295 | -91.4914
11.1299 | -91.4909
11.1302 | -91.4898
11.1305 | -91.489
11.1308 | -91.4884
11.1311 | -91.4879
11.1315 | -91.4873
11.1318 | -91.4868
11.1321 | -91.4863
11.1324 | -91.4857
11.1327 | -91.485
11.1331 | -91.4843
11.1334 | -91.4836
11.1337 | -91.4829
11.134 | -91.4822
11.1344 | -91.4812
11.1347 | -91.4802
11.135 | -91.4796
11.1353 | -91.4789
11.1356 | -91.4782
11.136 | -91.4774
11.1363 | -91.4768
11.1366 | -91.4767
11.1369 | -91.4762
11.1372 | -91.4756