## CALIBRATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 70 |  | CTBTO
Transducer Block | 02-90C/P | 67 |  | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
Hood | 02-92B/P | 50
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE REFERENCE
Atmospheric pressure | 1010.8 | hPa |  | PR 000 150_-_CAL
Humidity | 23.4 | %/Hr |  | Measurement device
Internal Temperature of Calibration Bench | 22.79 | °C |  | IS CALIBRATOR
AUX Sensor pressure | 1015 | hPa |  | Sensor SoH at test start
AUX Sensor Temperature | 23 | °C |  | OK
SENSITIVITY AT 1 Hz
Output Parameters | Theoretical Value | Measured Value @1Hz | Margin (±) | Unit | Result
Sensitivity of Pressure Derivative output | 2 | 2 | 0.04 | mV/Pa/sec | OK | Mode d'emploi pour créer un constat de vérification
Sensitivity of Pressure output | 20 | 19.88 | 0.4 | mV/Pa | OK | Cliquer sur "DEMARRER LA SAISIE"
Sensitivity of Calibration coil | 6 | 5.98 | 0.13 | Pa/V | OK | Suivre les instructions affichées
Une fois la partie SENSOR IDENTIFICATION remplie, renseigner la partie ENVIRONNEMENTAL PARAMETERS (les données environnementales)
TRANSFER FUNCTION | [5 empty] | Renseigner la partie SENSITIVITY AT 1Hz :les sensibilités en voie de dérivée de Pression (Pressure Derivative), en voie de pression (Pressure) et de Calibration (calibration Coil)
Electrical |  | Acoustical |  |  |  | Une sauvegarde temporaire est toujours possible si le fichier doit être modifié (Fichier "enregistrer")
Vérifier toutes les données saisie notamment la partie CUSTOMER ("Client") et la partie VISA (Noms des personnes et date d'ajustage)
Cliquer sur "ENREGISTREMENT DEFINITIF" (Attention le fichier sera enregistré et verrouillé dans le dossier courant, il faut retirer manuellement la protection)
SENSOR SELF NOISE |  |  |  | Decision : considering the specifications given by the manufacturer, this sensor is
CONFORM
VISA
|  |  |  | Adjusted by | A.HUE
|  |  |  | on | 2016-03-16
|  |  |  | Approved by | A.HUE
|  |  |  | Qualification | Optronic Engineer
Self Noise at 1 HZ | 5.9704e-05 | Pa/√Hz |  | on | 2016-03-17 | -84.48 | (dBPa/√Hz)
| < 0,13 | mPa/√Hz
OK
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which define the verification report. It can be used to demonstrate the link between the measure equipment and the standard national or international equipment, provided it meets the recommendations of the documentation booklet X 07-015
[8 empty] | Client | Adresse Client
[8 empty] | CEA | C.E.A. Dam. Bruyères Domaine du Grand Rué 91297 ARPAJON CEDEX FRANCE
[8 empty] | CTBTO | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
[8 empty] | IPGP | Université Paris Diderot Paris 7 Bat Lamarck UFR STEP 35 rue H.Brion 6ème étage 75205 Paris CEDEX 13
[8 empty] | SANDIA | C/O SANDIATIONAL LABS ALBUQUERQUE, NM 87123 UNITED STATES
[8 empty] | ISLA | Haiwaï Institute of Geophysics & Planetology 73-984 Makako Bay Drive KAILUA-KONA HAWAÏ
ENREGISTREMENT DEFINITIF
DEMARRER LA SAISIE

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.1403 | 11.7299
0.097656 | -34.0022 | 7.5771
0.12207 | -33.9067 | 5.3644
0.14648 | -33.9925 | 3.7513
0.1709 | -33.9796 | 3.3874
0.19531 | -33.9192 | 2.7224
0.21973 | -33.9567 | 2.3017
0.24414 | -34.0213 | 1.7245
0.26855 | -33.9965 | 2.1885
0.29297 | -33.9998 | 1.3984
0.31738 | -33.9046 | 1.0973
0.3418 | -33.9331 | 0.86866
0.36621 | -33.9182 | 1.0229
0.39062 | -33.9239 | 0.43977
0.41504 | -33.9557 | 0.4478
0.43945 | -33.9489 | 0.44278
0.46387 | -33.9003 | 0.11638
0.48828 | -33.9241 | 0.16342
0.5127 | -33.9935 | 0.19483
0.53711 | -33.9424 | -0.069874
0.56152 | -33.9672 | -0.2671
0.58594 | -33.9092 | -0.16005
0.61035 | -33.9635 | -0.44803
0.63477 | -33.9949 | -0.3374
0.65918 | -33.915 | -0.83458
0.68359 | -33.9161 | -0.76072
0.70801 | -33.9779 | -0.93429
0.73242 | -33.9721 | -0.47377
0.75684 | -34.0416 | -0.62628
0.78125 | -33.96 | -0.65266
0.80566 | -34.023 | -1.2272
0.83008 | -33.9752 | -0.90736
0.85449 | -33.8708 | -0.95408
0.87891 | -33.9514 | -1.2224
0.90332 | -33.9444 | -1.0777
0.92773 | -33.9593 | -1.123
0.95215 | -34.0119 | -1.2173
0.97656 | -33.9786 | -1.6068
1.001 | -33.9691 | -1.1107
1.0254 | -33.9892 | -1.268
1.0498 | -33.9772 | -1.1444
1.0742 | -34.0346 | -1.6793
1.0986 | -33.9573 | -1.7441
1.123 | -33.9467 | -1.4806
1.1475 | -33.9754 | -1.5199
1.1719 | -33.9779 | -1.8692
1.1963 | -33.965 | -1.9694
1.2207 | -33.9893 | -2.1329
1.2451 | -34.0023 | -1.7704
1.2695 | -33.9961 | -2.1974
1.2939 | -33.9898 | -1.6755
1.3184 | -34.0112 | -2.3816
1.3428 | -33.9995 | -1.9389
1.3672 | -34.0081 | -2.0966
1.3916 | -34.0063 | -2.3132
1.416 | -33.9563 | -2.0745
1.4404 | -33.9699 | -2.0715
1.4648 | -33.967 | -2.5502
1.4893 | -34.0491 | -2.2422
1.5137 | -34.0544 | -2.0592
1.5381 | -34.0657 | -2.6263
1.5625 | -34.0138 | -2.4148
1.5869 | -34.0032 | -2.6816
1.6113 | -34.0389 | -2.6865
1.6357 | -33.9787 | -2.9906
1.6602 | -33.9699 | -2.7423
1.6846 | -33.954 | -2.8791
1.709 | -34.0647 | -2.4062
1.7334 | -33.9368 | -2.8912
1.7578 | -34.0147 | -2.6099
1.7822 | -33.9695 | -2.5779
1.8066 | -34.0198 | -3.0663
1.8311 | -33.9601 | -2.8892
1.8555 | -34.0347 | -3.0864
1.8799 | -33.9852 | -3.3046
1.9043 | -33.952 | -3.1863
1.9287 | -34.0084 | -3.3293
1.9531 | -34.0127 | -3.1557
1.9775 | -33.972 | -3.1097
2.002 | -34.0283 | -3.8736
2.0264 | -34.0305 | -3.0949
2.0508 | -33.9302 | -3.3703
2.0752 | -33.9813 | -3.3732
2.0996 | -34.023 | -3.4419
2.124 | -34.0561 | -3.4864
2.1484 | -34.0026 | -3.6786
2.1729 | -34.0338 | -3.623
2.1973 | -33.999 | -3.9165
2.2217 | -34.0557 | -3.7699
2.2461 | -33.9857 | -3.9176
2.2705 | -34.0494 | -4.0321
2.2949 | -34.0263 | -4.006
2.3193 | -34.0342 | -3.744
2.3438 | -34.0398 | -4.2499
2.3682 | -34.0342 | -3.9333
2.3926 | -33.9819 | -4.0815
2.417 | -34.0126 | -3.9644
2.4414 | -33.9926 | -4.0324
2.4658 | -34.0149 | -4.23
2.4902 | -34.0469 | -4.0276
2.5146 | -34.0448 | -4.3301
2.5391 | -34.0327 | -4.4981
2.5635 | -34.064 | -4.329
2.5879 | -34.1205 | -4.3297
2.6123 | -34.0125 | -4.1037
2.6367 | -34.1007 | -4.6905
2.6611 | -34.0059 | -4.1866
2.6855 | -34.0691 | -4.855
2.71 | -34.017 | -4.3479
2.7344 | -34.0526 | -4.5857
2.7588 | -33.9827 | -4.7346
2.7832 | -34.0463 | -4.7334
2.8076 | -34.0134 | -4.648
2.832 | -34.0125 | -4.8412
2.8564 | -34.0642 | -4.811
2.8809 | -34.1251 | -5.2562
2.9053 | -34.018 | -4.8546
2.9297 | -33.9979 | -5.3889
2.9541 | -34.017 | -4.6782
2.9785 | -34.0847 | -5.2106
3.0029 | -34.0521 | -4.5833
3.0273 | -34.0366 | -5.3336
3.0518 | -34.0162 | -4.9866
3.0762 | -34.0605 | -5.4098
3.1006 | -34.1036 | -5.1202
3.125 | -34.0839 | -5.4079
3.1494 | -34.0201 | -5.2685
3.1738 | -34.1355 | -5.3565
3.1982 | -34.0887 | -5.9057
3.2227 | -34.0336 | -5.6764
3.2471 | -34.1001 | -5.6059
3.2715 | -34.018 | -5.8609
3.2959 | -33.9747 | -5.7484
3.3203 | -34.0404 | -5.469
3.3447 | -34.0501 | -5.7564
3.3691 | -34.0608 | -5.5651
3.3936 | -34.0997 | -5.7465
3.418 | -34.0869 | -5.7536
3.4424 | -34.0426 | -6.1375
3.4668 | -34.0886 | -6.0637
3.4912 | -34.0681 | -6.2031
3.5156 | -34.0884 | -6.061
3.54 | -34.0582 | -6.2389
3.5645 | -34.066 | -5.8268
3.5889 | -34.0486 | -6.0303
3.6133 | -34.1201 | -6.3364
3.6377 | -34.0678 | -6.3377
3.6621 | -34.0659 | -6.3196
3.6865 | -34.0591 | -6.2298
3.7109 | -34.0921 | -6.4132
3.7354 | -34.0446 | -6.2156
3.7598 | -34.0717 | -6.4699
3.7842 | -34.0996 | -6.1707
3.8086 | -34.0302 | -6.6771
3.833 | -34.0541 | -6.5711
3.8574 | -34.0583 | -6.7676
3.8818 | -34.0819 | -6.3263
3.9062 | -34.1167 | -7.2042
3.9307 | -34.1394 | -6.0212
3.9551 | -34.0779 | -7.1803
3.9795 | -34.0946 | -6.9691
4.0039 | -34.0918 | -6.5621
4.0283 | -34.1006 | -6.6412
4.0527 | -34.1231 | -7.2041
4.0771 | -34.0781 | -6.6447
4.1016 | -34.0999 | -6.9239
4.126 | -34.1623 | -7.1995
4.1504 | -34.0975 | -7.0373
4.1748 | -34.0774 | -6.9275
4.1992 | -34.0996 | -7.1605
4.2236 | -34.0405 | -7.5706
4.248 | -34.0459 | -7.5245
4.2725 | -34.1201 | -7.2524
4.2969 | -34.0885 | -7.169
4.3213 | -34.0909 | -7.4327
4.3457 | -34.113 | -7.0038
4.3701 | -34.0308 | -7.5391
4.3945 | -34.173 | -7.3279
4.4189 | -34.1097 | -7.364
4.4434 | -34.1396 | -7.0797
4.4678 | -34.1634 | -7.5782
4.4922 | -34.1442 | -8.1171
4.5166 | -34.0261 | -7.5101
4.541 | -34.1126 | -8.1444
4.5654 | -34.1396 | -7.7588
4.5898 | -34.1245 | -8.1968
4.6143 | -34.1414 | -7.7795
4.6387 | -34.0962 | -7.7942
4.6631 | -34.0962 | -7.6529
4.6875 | -34.1274 | -7.9372
4.7119 | -34.121 | -7.9307
4.7363 | -34.117 | -8.2687
4.7607 | -34.1393 | -8.2489
4.7852 | -34.1111 | -8.2637
4.8096 | -34.1975 | -8.0932
4.834 | -34.1111 | -8.6506
4.8584 | -34.1498 | -7.9677
4.8828 | -34.1205 | -8.0394
4.9072 | -34.1796 | -8.3438
4.9316 | -34.0754 | -8.4784
4.9561 | -34.1372 | -8.6487
4.9805 | -34.152 | -8.6702
5.0049 | -34.183 | -8.7207
5.0293 | -34.1201 | -8.3896
5.0537 | -34.1319 | -8.4531
5.0781 | -34.1236 | -8.5543
5.1025 | -34.1359 | -8.6998
5.127 | -34.116 | -8.7159
5.1514 | -34.1672 | -8.9375
5.1758 | -34.1551 | -8.9636
5.2002 | -34.1161 | -9.1028
5.2246 | -34.1651 | -9.0172
5.249 | -34.1007 | -9.2076
5.2734 | -34.1884 | -8.7104
5.2979 | -34.2101 | -8.8837
5.3223 | -34.1817 | -9.1205
5.3467 | -34.0821 | -9.1681
5.3711 | -34.1544 | -9.2393
5.3955 | -34.1683 | -9.3472
5.4199 | -34.1913 | -9.1824
5.4443 | -34.1258 | -9.1335
5.4688 | -34.1471 | -9.5985
5.4932 | -34.1425 | -9.5349
5.5176 | -34.156 | -9.4217
5.542 | -34.1317 | -9.3701
5.5664 | -34.1739 | -9.5782
5.5908 | -34.1228 | -9.755
5.6152 | -34.1012 | -9.6906
5.6396 | -34.1083 | -9.6442
5.6641 | -34.1176 | -9.8532
5.6885 | -34.1662 | -9.6328
5.7129 | -34.1505 | -9.3699
5.7373 | -34.1268 | -9.2156
5.7617 | -34.17 | -9.9335
5.7861 | -34.1419 | -9.5067
5.8105 | -34.152 | -9.7205
5.835 | -34.2137 | -9.4905
5.8594 | -34.1865 | -9.8729
5.8838 | -34.2069 | -10.2889
5.9082 | -34.1569 | -10.0077
5.9326 | -34.1537 | -10.3825
5.957 | -34.2149 | -10.265
5.9814 | -34.1706 | -10.4918
6.0059 | -34.1368 | -10.2384
6.0303 | -34.2015 | -10.418
6.0547 | -34.1765 | -10.2338
6.0791 | -34.156 | -10.383
6.1035 | -34.1951 | -10.0462
6.1279 | -34.1677 | -10.5967
6.1523 | -34.1639 | -10.6977
6.1768 | -34.1246 | -10.7461
6.2012 | -34.2098 | -10.6011
6.2256 | -34.1555 | -10.3618
6.25 | -34.2279 | -10.5353
6.2744 | -34.1574 | -10.3446
6.2988 | -34.2669 | -10.6182
6.3232 | -34.1391 | -10.7409
6.3477 | -34.1714 | -10.805
6.3721 | -34.1561 | -10.5471
6.3965 | -34.1864 | -10.7692
6.4209 | -34.2589 | -10.7085
6.4453 | -34.1774 | -11.2125
6.4697 | -34.1629 | -11.2208
6.4941 | -34.097 | -11.1882
6.5186 | -34.2942 | -11.2147
6.543 | -34.1832 | -11.4465
6.5674 | -34.1493 | -10.7849
6.5918 | -34.2457 | -11.2663
6.6162 | -34.1796 | -10.9496
6.6406 | -34.1689 | -10.904
6.665 | -34.2651 | -11.8663
6.6895 | -34.2113 | -12.0094
6.7139 | -34.1729 | -11.3693
6.7383 | -34.167 | -10.932
6.7627 | -34.2283 | -11.9267
6.7871 | -34.2251 | -11.5435
6.8115 | -34.1709 | -11.0206
6.8359 | -34.2041 | -10.6143
6.8604 | -34.3047 | -11.3218
6.8848 | -34.2571 | -11.4244
6.9092 | -34.1963 | -11.5531
6.9336 | -34.3527 | -12.0694
6.958 | -34.2262 | -12.4354
6.9824 | -34.1389 | -11.5814
7.0068 | -34.1463 | -11.754
7.0312 | -34.2463 | -12.0528
7.0557 | -34.2488 | -11.7555
7.0801 | -34.3246 | -11.6864
7.1045 | -34.215 | -11.5887
7.1289 | -34.3346 | -12.3553
7.1533 | -34.2791 | -12.6716
7.1777 | -34.2462 | -11.411
7.2021 | -34.2043 | -11.9985
7.2266 | -34.1967 | -12.4987
7.251 | -34.208 | -11.6635
7.2754 | -34.3527 | -12.2843
7.2998 | -34.1752 | -11.9007
7.3242 | -34.1544 | -11.8962
7.3486 | -34.2875 | -11.844
7.373 | -34.1963 | -13.2328
7.3975 | -34.3521 | -12.4615
7.4219 | -34.2143 | -12.9758
7.4463 | -34.3135 | -12.0104
7.4707 | -34.2494 | -13.1815
7.4951 | -34.2959 | -12.766
7.5195 | -34.2514 | -12.7713
7.5439 | -34.324 | -12.426
7.5684 | -34.253 | -12.5346
7.5928 | -34.411 | -12.8341
7.6172 | -34.2299 | -12.6841
7.6416 | -34.3131 | -12.9391
7.666 | -34.1808 | -12.727
7.6904 | -34.2266 | -12.8632
7.7148 | -34.3206 | -12.814
7.7393 | -34.3294 | -12.8733
7.7637 | -34.2497 | -12.9659
7.7881 | -34.2986 | -13.3892
7.8125 | -34.2587 | -13.4932
7.8369 | -34.2974 | -13.6112
7.8613 | -34.2353 | -13.1187
7.8857 | -34.2468 | -13.2838
7.9102 | -34.2944 | -13.1153
7.9346 | -34.2793 | -13.2698
7.959 | -34.3544 | -13.3466
7.9834 | -34.27 | -12.5894
8.0078 | -34.3993 | -13.3128
8.0322 | -34.3199 | -13.251
8.0566 | -34.2535 | -12.892
8.0811 | -34.2496 | -13.0985
8.1055 | -34.3411 | -13.9092
8.1299 | -34.2198 | -13.1605
8.1543 | -34.2497 | -13.6739
8.1787 | -34.4254 | -13.3494
8.2031 | -34.3665 | -13.6491
8.2275 | -34.2463 | -13.617
8.252 | -34.3926 | -13.6054
8.2764 | -34.4089 | -14.734
8.3008 | -34.348 | -14.4546
8.3252 | -34.2635 | -14.0448
8.3496 | -34.2538 | -13.4986
8.374 | -34.3952 | -13.3311
8.3984 | -34.343 | -14.6504
8.4229 | -34.2406 | -13.5258
8.4473 | -34.2892 | -13.9014
8.4717 | -34.3688 | -13.8004
8.4961 | -34.3635 | -14.2124
8.5205 | -34.3219 | -13.757
8.5449 | -34.3706 | -13.8374
8.5693 | -34.4123 | -13.7394
8.5938 | -34.3963 | -13.928
8.6182 | -34.3606 | -15.4167
8.6426 | -34.1555 | -14.4602
8.667 | -34.3763 | -14.5661
8.6914 | -34.2337 | -14.2122
8.7158 | -34.3245 | -14.4152
8.7402 | -34.3229 | -13.9347
8.7646 | -34.3196 | -14.2265
8.7891 | -34.4334 | -14.3329
8.8135 | -34.4384 | -15.0393
8.8379 | -34.2515 | -14.3098
8.8623 | -34.4754 | -14.6298
8.8867 | -34.4442 | -15.7239
8.9111 | -34.2657 | -15.6406
8.9355 | -34.2938 | -14.4158
8.96 | -34.3446 | -14.7503
8.9844 | -34.3508 | -15.6394
9.0088 | -34.297 | -14.669
9.0332 | -34.4544 | -15.1912
9.0576 | -34.235 | -15.0235
9.082 | -34.382 | -14.8141
9.1064 | -34.2889 | -14.941
9.1309 | -34.4345 | -14.8648
9.1553 | -34.3648 | -15.6251
9.1797 | -34.3694 | -15.3317
9.2041 | -34.3953 | -15.8413
9.2285 | -34.2472 | -15.8138
9.2529 | -34.4022 | -15.5988
9.2773 | -34.2665 | -15.0799
9.3018 | -34.335 | -15.5255
9.3262 | -34.3927 | -14.6649
9.3506 | -34.4744 | -15.2241
9.375 | -34.4994 | -15.398
9.3994 | -34.4935 | -15.3701
9.4238 | -34.3047 | -16.1727
9.4482 | -34.4595 | -14.8259
9.4727 | -34.4639 | -15.7712
9.4971 | -34.2728 | -15.8763
9.5215 | -34.4216 | -15.3422
9.5459 | -34.3713 | -16.5495
9.5703 | -34.3778 | -14.899
9.5947 | -34.4631 | -16.0486
9.6191 | -34.3069 | -15.5009
9.6436 | -34.3933 | -15.288
9.668 | -34.5335 | -16.148
9.6924 | -34.2751 | -15.7646
9.7168 | -34.5149 | -16.0424
9.7412 | -34.4791 | -15.9207
9.7656 | -34.4815 | -15.8506
9.79 | -34.2617 | -16.2479
9.8145 | -34.4028 | -15.4279
9.8389 | -34.5307 | -16.7864
9.8633 | -34.4449 | -16.7787
9.8877 | -34.3025 | -15.8319
9.9121 | -34.2752 | -15.9825
9.9365 | -34.4883 | -16.1793
9.9609 | -34.4098 | -16.0727
9.9854 | -34.3318 | -15.8718
10.0098 | -34.5037 | -16.6455

## ACOUS_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.1511 | 9.5912
0.097656 | -34.1946 | 5.079
0.12207 | -34.0531 | 3.5648
0.14648 | -34.1408 | 1.5045
0.1709 | -34.065 | 1.2624
0.19531 | -34.0845 | 0.15313
0.21973 | -34.0855 | 0.39276
0.24414 | -34.0977 | -0.81606
0.26855 | -34.1044 | -0.31823
0.29297 | -34.0521 | -1.4592
0.31738 | -34.0989 | -0.9964
0.3418 | -34.0371 | -1.725
0.36621 | -34.1199 | -1.3286
0.39062 | -34.0266 | -2.0556
0.41504 | -34.1409 | -1.7104
0.43945 | -34.0457 | -2.1737
0.46387 | -34.1443 | -2.2967
0.48828 | -34.008 | -2.243
0.5127 | -34.1669 | -2.588
0.53711 | -34.0489 | -2.3025
0.56152 | -34.1424 | -2.8991
0.58594 | -34.0608 | -2.45
0.61035 | -34.1325 | -3.2433
0.63477 | -34.0858 | -2.5589
0.65918 | -34.1237 | -3.4358
0.68359 | -34.0808 | -2.6726
0.70801 | -34.0871 | -3.5267
0.73242 | -34.1264 | -2.8798
0.75684 | -34.0909 | -3.7042
0.78125 | -34.1449 | -3.0367
0.80566 | -34.0662 | -3.8023
0.83008 | -34.1642 | -3.3687
0.85449 | -34.0402 | -3.8067
0.87891 | -34.1304 | -3.5457
0.90332 | -34.027 | -3.5728
0.92773 | -34.1756 | -3.612
0.95215 | -34.0743 | -3.6049
0.97656 | -34.1783 | -4.0767
1.001 | -34.0726 | -3.7048
1.0254 | -34.1751 | -4.2551
1.0498 | -34.1122 | -3.7386
1.0742 | -34.1671 | -4.5955
1.0986 | -34.091 | -3.9281
1.123 | -34.1182 | -4.6216
1.1475 | -34.1209 | -3.859
1.1719 | -34.115 | -4.7537
1.1963 | -34.1449 | -3.9014
1.2207 | -34.1028 | -4.8656
1.2451 | -34.1808 | -4.1154
1.2695 | -34.0791 | -4.7826
1.2939 | -34.1813 | -4.2946
1.3184 | -34.0517 | -4.6549
1.3428 | -34.2093 | -4.4489
1.3672 | -34.0899 | -4.6642
1.3916 | -34.1921 | -4.8012
1.416 | -34.0521 | -4.7636
1.4404 | -34.1818 | -4.9603
1.4648 | -34.0664 | -4.4269
1.4893 | -34.1858 | -4.9822
1.5137 | -34.115 | -4.4661
1.5381 | -34.1833 | -5.3907
1.5625 | -34.1455 | -4.5865
1.5869 | -34.1673 | -5.4603
1.6113 | -34.1674 | -4.7877
1.6357 | -34.1133 | -5.7275
1.6602 | -34.1592 | -4.9702
1.6846 | -34.0972 | -5.531
1.709 | -34.1817 | -5.023
1.7334 | -34.0708 | -5.5799
1.7578 | -34.2131 | -5.3022
1.7822 | -34.0721 | -5.4601
1.8066 | -34.2059 | -5.3746
1.8311 | -34.0808 | -5.3652
1.8555 | -34.2347 | -5.5845
1.8799 | -34.1066 | -5.3641
1.9043 | -34.1989 | -5.9797
1.9287 | -34.0884 | -5.4284
1.9531 | -34.1836 | -6.0554
1.9775 | -34.1235 | -5.2033
2.002 | -34.1732 | -6.1722
2.0264 | -34.1408 | -5.3705
2.0508 | -34.1518 | -6.1115
2.0752 | -34.1795 | -5.4468
2.0996 | -34.157 | -6.2727
2.124 | -34.2085 | -5.548
2.1484 | -34.1359 | -6.4909
2.1729 | -34.2082 | -5.9113
2.1973 | -34.1196 | -6.2615
2.2217 | -34.2189 | -6.0135
2.2461 | -34.1152 | -6.2504
2.2705 | -34.2562 | -6.2198
2.2949 | -34.1261 | -6.27
2.3193 | -34.2278 | -6.4359
2.3438 | -34.1262 | -6.1371
2.3682 | -34.1946 | -6.6293
2.3926 | -34.1213 | -5.983
2.417 | -34.1854 | -6.7285
2.4414 | -34.1573 | -5.9316
2.4658 | -34.2173 | -6.7937
2.4902 | -34.2034 | -6.0892
2.5146 | -34.1776 | -7.0505
2.5391 | -34.1938 | -6.2978
2.5635 | -34.1618 | -6.9891
2.5879 | -34.2179 | -6.4073
2.6123 | -34.135 | -6.8923
2.6367 | -34.2274 | -6.5522
2.6611 | -34.1342 | -6.9982
2.6855 | -34.2567 | -6.6689
2.71 | -34.1189 | -6.9069
2.7344 | -34.2615 | -6.9155
2.7588 | -34.1456 | -6.6641
2.7832 | -34.2391 | -7.2339
2.8076 | -34.1312 | -6.6928
2.832 | -34.2394 | -7.2687
2.8564 | -34.1833 | -6.8056
2.8809 | -34.2145 | -7.6558
2.9053 | -34.1693 | -6.7938
2.9297 | -34.2051 | -7.6253
2.9541 | -34.2213 | -6.9339
2.9785 | -34.1672 | -7.7055
3.0029 | -34.227 | -7.011
3.0273 | -34.1638 | -7.7007
3.0518 | -34.2585 | -7.0651
3.0762 | -34.16 | -7.586
3.1006 | -34.262 | -7.2665
3.125 | -34.1551 | -7.5686
3.1494 | -34.2593 | -7.4765
3.1738 | -34.1593 | -7.6501
3.1982 | -34.2682 | -7.7379
3.2227 | -34.1345 | -7.5195
3.2471 | -34.2527 | -7.8768
3.2715 | -34.1681 | -7.4122
3.2959 | -34.2486 | -7.8664
3.3203 | -34.1838 | -7.2613
3.3447 | -34.2216 | -8.1424
3.3691 | -34.2175 | -7.4632
3.3936 | -34.2223 | -8.1193
3.418 | -34.2282 | -7.5071
3.4424 | -34.205 | -8.3335
3.4668 | -34.2659 | -7.6246
3.4912 | -34.1855 | -8.428
3.5156 | -34.3036 | -7.9265
3.54 | -34.1523 | -8.358
3.5645 | -34.2821 | -8.0733
3.5889 | -34.1971 | -8.2269
3.6133 | -34.3033 | -8.11
3.6377 | -34.1616 | -8.2866
3.6621 | -34.2863 | -8.4358
3.6865 | -34.1737 | -8.0322
3.7109 | -34.3005 | -8.4626
3.7354 | -34.214 | -8.1314
3.7598 | -34.2741 | -8.76
3.7842 | -34.2165 | -8.2314
3.8086 | -34.2389 | -8.854
3.833 | -34.227 | -8.1742
3.8574 | -34.2368 | -8.7949
3.8818 | -34.2828 | -8.2082
3.9062 | -34.2205 | -9.0421
3.9307 | -34.2688 | -8.2418
3.9551 | -34.2121 | -8.9348
3.9795 | -34.3296 | -8.5286
4.0039 | -34.1994 | -8.8953
4.0283 | -34.3122 | -8.7198
4.0527 | -34.1732 | -8.8664
4.0771 | -34.3144 | -8.9684
4.1016 | -34.2183 | -8.8011
4.126 | -34.3225 | -9.069
4.1504 | -34.2129 | -8.838
4.1748 | -34.3045 | -9.404
4.1992 | -34.2306 | -8.8599
4.2236 | -34.2644 | -9.6014
4.248 | -34.2295 | -8.7519
4.2725 | -34.2431 | -9.4739
4.2969 | -34.2499 | -8.6643
4.3213 | -34.2667 | -9.5732
4.3457 | -34.2729 | -8.9506
4.3701 | -34.2235 | -9.6199
4.3945 | -34.2992 | -8.9878
4.4189 | -34.2197 | -9.541
4.4434 | -34.3415 | -9.0814
4.4678 | -34.2475 | -9.5867
4.4922 | -34.3338 | -9.5514
4.5166 | -34.1857 | -9.507
4.541 | -34.331 | -9.6676
4.5654 | -34.2126 | -9.4051
4.5898 | -34.3288 | -9.9098
4.6143 | -34.2238 | -9.3313
4.6387 | -34.3089 | -10.0363
4.6631 | -34.2584 | -9.3483
4.6875 | -34.2979 | -10.1685
4.7119 | -34.2668 | -9.4323
4.7363 | -34.2811 | -9.9099
4.7607 | -34.2849 | -9.4178
4.7852 | -34.2813 | -10.2593
4.8096 | -34.3218 | -9.6517
4.834 | -34.2405 | -10.3037
4.8584 | -34.3536 | -9.604
4.8828 | -34.2291 | -9.9977
4.9072 | -34.3625 | -9.8246
4.9316 | -34.2342 | -10.2486
4.9561 | -34.3624 | -10.2248
4.9805 | -34.2262 | -9.9882
5.0049 | -34.3363 | -10.4395
5.0293 | -34.2042 | -10.0263
5.0537 | -34.322 | -10.5446
5.0781 | -34.2392 | -9.8003
5.1025 | -34.3403 | -10.6704
5.127 | -34.2669 | -9.9695
5.1514 | -34.3306 | -10.7834
5.1758 | -34.3065 | -10.0314
5.2002 | -34.3188 | -10.8522
5.2246 | -34.3289 | -10.0177
5.249 | -34.2555 | -10.8265
5.2734 | -34.3327 | -10.0695
5.2979 | -34.2811 | -10.9473
5.3223 | -34.3858 | -10.4307
5.3467 | -34.251 | -10.7603
5.3711 | -34.3883 | -10.5628
5.3955 | -34.2309 | -10.7713
5.4199 | -34.4033 | -10.8737
5.4443 | -34.2655 | -10.6904
5.4688 | -34.354 | -11.2172
5.4932 | -34.2388 | -10.6123
5.5176 | -34.3524 | -11.1337
5.542 | -34.2829 | -10.6863
5.5664 | -34.3508 | -11.3442
5.5908 | -34.2971 | -10.6954
5.6152 | -34.3328 | -11.4859
5.6396 | -34.3056 | -10.8764
5.6641 | -34.294 | -11.5472
5.6885 | -34.3448 | -10.7794
5.7129 | -34.2699 | -11.5402
5.7373 | -34.3473 | -10.9308
5.7617 | -34.2945 | -11.388
5.7861 | -34.414 | -10.8069
5.8105 | -34.2662 | -11.4293
5.835 | -34.4265 | -11.2383
5.8594 | -34.2658 | -11.3797
5.8838 | -34.4092 | -11.4335
5.9082 | -34.2797 | -11.3678
5.9326 | -34.3961 | -11.7742
5.957 | -34.2976 | -11.0787
5.9814 | -34.391 | -11.8149
6.0059 | -34.3118 | -11.3097
6.0303 | -34.3616 | -11.9694
6.0547 | -34.3175 | -11.3131
6.0791 | -34.3252 | -11.8436
6.1035 | -34.3503 | -11.065
6.1279 | -34.3285 | -12.1567
6.1523 | -34.3738 | -11.4673
6.1768 | -34.294 | -12.2261
6.2012 | -34.3802 | -11.6906
6.2256 | -34.2939 | -11.9508
6.25 | -34.417 | -11.6308
6.2744 | -34.3032 | -11.8419
6.2988 | -34.4012 | -12.0082
6.3232 | -34.3001 | -11.7681
6.3477 | -34.4212 | -11.8958
6.3721 | -34.2897 | -11.6771
6.3965 | -34.4083 | -12.2308
6.4209 | -34.3122 | -11.7912
6.4453 | -34.3915 | -12.3337
6.4697 | -34.3401 | -11.6808
6.4941 | -34.3852 | -12.7022
6.5186 | -34.386 | -11.78
6.543 | -34.377 | -12.7444
6.5674 | -34.4116 | -11.8727
6.5918 | -34.3536 | -12.8347
6.6162 | -34.3995 | -12.1247
6.6406 | -34.2977 | -12.6452
6.665 | -34.4418 | -12.201
6.6895 | -34.3118 | -12.4496
6.7139 | -34.4435 | -12.4717
6.7383 | -34.3041 | -12.5555
6.7627 | -34.4202 | -12.6616
6.7871 | -34.2876 | -12.3773
6.8115 | -34.4604 | -12.6901
6.8359 | -34.3516 | -12.628
6.8604 | -34.4348 | -12.9437
6.8848 | -34.3221 | -12.3811
6.9092 | -34.4151 | -13.0494
6.9336 | -34.404 | -12.5018
6.958 | -34.3958 | -13.2686
6.9824 | -34.4076 | -12.4251
7.0068 | -34.3905 | -13.1806
7.0312 | -34.4169 | -12.4119
7.0557 | -34.3758 | -13.1248
7.0801 | -34.4387 | -12.4714
7.1045 | -34.3843 | -13.0438
7.1289 | -34.481 | -12.7794
7.1533 | -34.3448 | -13.1532
7.1777 | -34.4686 | -12.953
7.2021 | -34.3694 | -13.1619
7.2266 | -34.504 | -13.0962
7.251 | -34.3971 | -13.1793
7.2754 | -34.4695 | -13.3129
7.2998 | -34.3416 | -12.7546
7.3242 | -34.4384 | -13.5544
7.3486 | -34.3941 | -12.8282
7.373 | -34.4427 | -13.6069
7.3975 | -34.4196 | -12.8364
7.4219 | -34.4372 | -13.8113
7.4463 | -34.4261 | -12.9985
7.4707 | -34.3897 | -13.7195
7.4951 | -34.4128 | -13.1675
7.5195 | -34.3824 | -13.6516
7.5439 | -34.4655 | -13.263
7.5684 | -34.3669 | -13.5521
7.5928 | -34.4439 | -13.383
7.6172 | -34.3424 | -13.5559
7.6416 | -34.5071 | -13.5481
7.666 | -34.3586 | -13.3258
7.6904 | -34.5265 | -13.8439
7.7148 | -34.4076 | -13.5364
7.7393 | -34.495 | -13.9991
7.7637 | -34.4527 | -13.5381
7.7881 | -34.4932 | -14.3148
7.8125 | -34.4313 | -13.5956
7.8369 | -34.4352 | -14.3327
7.8613 | -34.4411 | -13.428
7.8857 | -34.4312 | -14.513
7.9102 | -34.4869 | -13.5586
7.9346 | -34.3985 | -14.3915
7.959 | -34.5003 | -13.7686
7.9834 | -34.4071 | -14.3654
8.0078 | -34.5206 | -13.8622
8.0322 | -34.4154 | -14.3428
8.0566 | -34.5342 | -14.1404
8.0811 | -34.4085 | -14.0686
8.1055 | -34.5207 | -14.1051
8.1299 | -34.3863 | -14.0544
8.1543 | -34.5032 | -14.5149
8.1787 | -34.4143 | -14.0334
8.2031 | -34.4659 | -14.4127
8.2275 | -34.4439 | -14.0169
8.252 | -34.4871 | -14.7489
8.2764 | -34.4744 | -14.2106
8.3008 | -34.4808 | -14.7512
8.3252 | -34.4863 | -14.1956
8.3496 | -34.4727 | -14.9398
8.374 | -34.5292 | -14.0897
8.3984 | -34.4441 | -15.1393
8.4229 | -34.5047 | -14.3082
8.4473 | -34.4536 | -14.9125
8.4717 | -34.525 | -14.4086
8.4961 | -34.4148 | -14.9942
8.5205 | -34.5516 | -14.3735
8.5449 | -34.4169 | -14.8791
8.5693 | -34.5814 | -14.7523
8.5938 | -34.4274 | -14.4528
8.6182 | -34.5612 | -15.1789
8.6426 | -34.446 | -14.4928
8.667 | -34.5184 | -15.2431
8.6914 | -34.4527 | -14.2589
8.7158 | -34.4923 | -15.4528
8.7402 | -34.4652 | -14.5285
8.7646 | -34.4934 | -15.2694
8.7891 | -34.5329 | -14.7164
8.8135 | -34.4877 | -15.6046
8.8379 | -34.5676 | -14.5708
8.8623 | -34.4798 | -15.2324
8.8867 | -34.5827 | -15.0352
8.9111 | -34.4572 | -15.1614
8.9355 | -34.5676 | -15.2464
8.96 | -34.4645 | -15.4133
8.9844 | -34.5783 | -15.4358
9.0088 | -34.4663 | -15.3322
9.0332 | -34.5824 | -15.4247
9.0576 | -34.427 | -15.1874
9.082 | -34.603 | -15.647
9.1064 | -34.5021 | -14.9052
9.1309 | -34.5691 | -15.7416
9.1553 | -34.488 | -15.0906
9.1797 | -34.5292 | -15.9526
9.2041 | -34.5243 | -15.2446
9.2285 | -34.5305 | -15.9288
9.2529 | -34.5122 | -15.0301
9.2773 | -34.481 | -16.1605
9.3018 | -34.5616 | -15.1851
9.3262 | -34.5123 | -15.8823
9.3506 | -34.6167 | -15.4852
9.375 | -34.4768 | -15.6922
9.3994 | -34.5912 | -15.4109
9.4238 | -34.5042 | -15.673
9.4482 | -34.6387 | -15.5286
9.4727 | -34.4969 | -15.7606
9.4971 | -34.5922 | -16.0735
9.5215 | -34.4753 | -15.6228
9.5459 | -34.5745 | -16.3393
9.5703 | -34.5174 | -15.4798
9.5947 | -34.5555 | -16.4527
9.6191 | -34.5274 | -15.6871
9.6436 | -34.5821 | -16.3934
9.668 | -34.5626 | -15.6465
9.6924 | -34.503 | -16.352
9.7168 | -34.5751 | -15.7593
9.7412 | -34.5241 | -16.4793
9.7656 | -34.6144 | -15.7108
9.79 | -34.4909 | -16.4569
9.8145 | -34.6218 | -15.8173
9.8389 | -34.498 | -16.7529
9.8633 | -34.6318 | -16.1466
9.8877 | -34.4826 | -16.2722
9.9121 | -34.6216 | -16.2265
9.9365 | -34.5072 | -16.3781
9.9609 | -34.6397 | -16.4055
9.9854 | -34.5322 | -16.213
10.0098 | -34.6564 | -16.9218

## SEALED_NOISE
dBPa/√Hz | Frequency
0.009316 | -37.3866
0.009627 | -37.7875
0.009937 | -37.9194
0.010248 | -34.5676
0.010558 | -32.4775
0.010869 | -32.2591
0.011179 | -32.5099
0.01149 | -34.6289
0.0118 | -35.588
0.012111 | -36.2184
0.012422 | -37.0893
0.012732 | -36.1893
0.013043 | -36.1164
0.013353 | -35.352
0.013664 | -34.7182
0.013974 | -34.8988
0.014285 | -33.5414
0.014595 | -33.7123
0.014906 | -35.9953
0.015216 | -35.3215
0.015527 | -32.9006
0.015837 | -31.5288
0.016148 | -30.9934
0.016459 | -31.2309
0.016769 | -30.3521
0.01739 | -31.8381
0.017701 | -34.5047
0.018011 | -37.9935
0.018322 | -41.2275
0.018632 | -41.6775
0.018943 | -39.8535
0.019253 | -38.2838
0.019564 | -37.7041
0.019874 | -37.8249
0.020185 | -38.3969
0.020495 | -39.4224
0.020806 | -40.6545
0.021117 | -41.6648
0.021427 | -41.4782
0.021738 | -39.8376
0.022048 | -38.3282
0.022359 | -37.9092
0.022669 | -38.2547
0.02298 | -38.0425
0.02329 | -37.195
0.023601 | -36.9749
0.023911 | -37.711
0.024222 | -38.8936
0.024532 | -39.5747
0.024843 | -39.6234
0.025154 | -39.4154
0.025464 | -38.5553
0.025775 | -37.2688
0.026085 | -36.7734
0.026396 | -37.639
0.026706 | -39.5098
0.027017 | -41.5982
0.027327 | -43.8888
0.027638 | -45.6982
0.027948 | -46.3161
0.028259 | -46.0808
0.028569 | -45.4088
0.02888 | -44.559
0.029191 | -43.628
0.029501 | -42.8344
0.029812 | -42.4431
0.030122 | -42.4856
0.030433 | -42.5861
0.030743 | -42.3193
0.031054 | -41.5534
0.031364 | -40.7831
0.031675 | -40.5741
0.031985 | -41.2292
0.032296 | -42.5863
0.032606 | -43.6898
0.032917 | -43.1366
0.033228 | -41.4363
0.033538 | -39.9825
0.033849 | -39.3907
0.034159 | -39.7377
0.03447 | -40.8247
0.03478 | -42.1293
0.035091 | -43.1076
0.035401 | -43.6949
0.035712 | -43.8755
0.036022 | -43.9758
0.036333 | -43.8184
0.036643 | -43.5664
0.036954 | -43.517
0.037265 | -43.796
0.037575 | -44.2977
0.037886 | -44.8453
0.038196 | -44.8574
0.038507 | -44.22
0.038817 | -43.4371
0.039128 | -42.9092
0.039438 | -42.9543
0.039749 | -43.541
0.040059 | -44.5485
0.04037 | -45.6811
0.04068 | -46.5428
0.040991 | -46.9992
0.041302 | -46.8783
0.041612 | -46.6793
0.041923 | -46.5946
0.042233 | -46.7784
0.042544 | -47.2918
0.042854 | -48.0312
0.043165 | -48.8459
0.043475 | -49.5674
0.043786 | -49.9324
0.044096 | -50.1174
0.044407 | -49.8923
0.044717 | -49.1707
0.045028 | -48.7127
0.045339 | -48.387
0.045649 | -48.3033
0.04596 | -48.4292
0.04627 | -48.635
0.046581 | -48.8051
0.046891 | -48.9109
0.047202 | -48.9527
0.047512 | -49.0746
0.047823 | -49.1431
0.048133 | -49.2103
0.048444 | -49.3005
0.048754 | -49.4624
0.049065 | -49.8238
0.049376 | -50.3789
0.049686 | -51.0644
0.049997 | -51.7447
0.050307 | -52.265
0.050618 | -52.3834
0.050928 | -51.5522
0.051239 | -50.3186
0.051549 | -48.8026
0.05186 | -47.3159
0.05217 | -46.0747
0.052481 | -45.2367
0.052791 | -44.8476
0.053102 | -44.8704
0.053412 | -45.1756
0.053723 | -45.5911
0.054034 | -45.8816
0.054344 | -46.0893
0.054655 | -46.129
0.054965 | -46.0637
0.055276 | -46.0087
0.055586 | -45.8823
0.055897 | -45.7806
0.056207 | -45.7918
0.056518 | -45.889
0.056828 | -46.0543
0.057139 | -46.3136
0.057449 | -46.6246
0.05776 | -46.9247
0.058071 | -47.0998
0.058381 | -47.1874
0.058692 | -47.3182
0.059002 | -47.5389
0.059313 | -47.9716
0.059623 | -48.5121
0.059934 | -49.0326
0.060244 | -49.4334
0.060555 | -49.6196
0.060865 | -49.5084
0.061176 | -49.2704
0.061486 | -48.9791
0.061797 | -48.7132
0.062108 | -48.627
0.062418 | -48.9991
0.062729 | -49.2501
0.063039 | -49.5464
0.06335 | -49.8147
0.06366 | -49.9725
0.063971 | -49.9443
0.064281 | -49.6779
0.064592 | -49.31
0.064902 | -48.9509
0.065213 | -48.6726
0.065523 | -48.5272
0.065834 | -48.5899
0.066145 | -48.8761
0.066455 | -49.3734
0.066766 | -50.0606
0.067076 | -50.8825
0.067387 | -51.7096
0.067697 | -52.4205
0.068008 | -52.9755
0.068318 | -53.3317
0.068629 | -53.4396
0.068939 | -53.4185
0.06925 | -53.3827
0.06956 | -53.4548
0.069871 | -53.6591
0.070182 | -54.0487
0.070492 | -54.5539
0.070803 | -55.1436
0.071113 | -55.7359
0.071424 | -55.8628
0.071734 | -56.131
0.072045 | -55.9628
0.072355 | -55.4967
0.072666 | -54.9274
0.072976 | -54.2802
0.073287 | -53.5172
0.073597 | -52.8294
0.073908 | -52.2351
0.074219 | -51.7633
0.074529 | -51.432
0.07484 | -51.2426
0.07515 | -51.2036
0.075461 | -51.2967
0.075771 | -51.4852
0.076082 | -51.7292
0.076392 | -51.9694
0.076703 | -52.2029
0.077013 | -52.4411
0.077324 | -52.5934
0.077634 | -52.7416
0.077945 | -52.9146
0.078256 | -53.031
0.078566 | -53.1177
0.078877 | -53.2404
0.079187 | -53.414
0.079498 | -53.5691
0.079808 | -53.7274
0.080119 | -53.9423
0.080429 | -54.0648
0.08074 | -54.2165
0.08105 | -54.3524
0.081361 | -54.4279
0.081671 | -54.4658
0.081982 | -54.4707
0.082293 | -54.4829
0.082603 | -54.5267
0.082914 | -54.5615
0.083224 | -54.6295
0.083535 | -54.8104
0.083845 | -55.0799
0.084156 | -55.4728
0.084466 | -55.8605
0.084777 | -56.2875
0.085087 | -56.7166
0.085398 | -57.0865
0.085708 | -57.3754
0.086019 | -57.6143
0.08633 | -57.6766
0.08664 | -57.7219
0.086951 | -57.7292
0.087261 | -57.5904
0.087572 | -57.4225
0.087882 | -57.3096
0.088193 | -57.2039
0.088503 | -57.1003
0.088814 | -57.0143
0.089124 | -56.9104
0.089435 | -56.8474
0.089745 | -56.8137
0.090056 | -56.8099
0.090366 | -56.7776
0.090677 | -56.734
0.090988 | -56.7335
0.091298 | -56.7763
0.091609 | -56.7463
0.091919 | -56.6916
0.09223 | -56.619
0.09254 | -56.518
0.092851 | -56.3804
0.093161 | -56.265
0.093472 | -56.1786
0.093782 | -56.1575
0.094093 | -56.2343
0.094403 | -56.3882
0.094714 | -56.5977
0.095025 | -56.8936
0.095335 | -57.2711
0.095646 | -57.7177
0.095956 | -58.2302
0.096267 | -58.7321
0.096577 | -59.1704
0.096888 | -59.556
0.097198 | -59.8653
0.097509 | -60.1478
0.097819 | -60.3523
0.09813 | -60.4388
0.09844 | -60.3509
0.098751 | -60.2761
0.099062 | -60.2038
0.099372 | -60.1326
0.099683 | -59.961
0.099993 | -59.7678
0.1003 | -59.5127
0.10061 | -59.244
0.10093 | -58.9673
0.10124 | -58.7016
0.10155 | -58.4507
0.10186 | -58.2161
0.10217 | -58.0456
0.10248 | -57.9549
0.10279 | -57.9283
0.1031 | -57.9407
0.10341 | -58.0086
0.10372 | -58.1398
0.10403 | -58.3082
0.10434 | -58.4754
0.10465 | -58.6463
0.10496 | -58.8239
0.10527 | -59.0245
0.10558 | -59.1804
0.10589 | -59.3399
0.1062 | -59.4485
0.10651 | -59.5065
0.10683 | -59.5196
0.10714 | -59.4549
0.10745 | -59.4007
0.10776 | -59.3798
0.10807 | -59.377
0.10838 | -59.4084
0.10869 | -59.4543
0.109 | -59.5465
0.10931 | -59.7016
0.10962 | -59.9032
0.10993 | -60.1812
0.11024 | -60.4578
0.11055 | -60.7168
0.11086 | -60.967
0.11117 | -61.1908
0.11148 | -61.3588
0.11179 | -61.4664
0.1121 | -61.4896
0.11242 | -61.4672
0.11273 | -61.4002
0.11304 | -61.2598
0.11335 | -61.0756
0.11366 | -60.8995
0.11397 | -60.7386
0.11428 | -60.6091
0.11459 | -60.4984
0.1149 | -60.3974
0.11521 | -60.3835
0.11552 | -60.3841
0.11583 | -60.4079
0.11614 | -60.4779
0.11645 | -60.5931
0.11676 | -60.6665
0.11707 | -60.7458
0.11738 | -60.8439
0.11769 | -60.9286
0.118 | -60.9891
0.11832 | -61.0734
0.11862 | -61.1616
0.11894 | -61.2466
0.11925 | -61.3616
0.11956 | -61.5022
0.11987 | -61.6676
0.12018 | -61.8557
0.12049 | -62.0538
0.1208 | -62.272
0.12111 | -62.5199
0.12142 | -62.7275
0.12173 | -62.9479
0.12204 | -63.1786
0.12235 | -63.4176
0.12266 | -63.6365
0.12297 | -63.8114
0.12328 | -63.9873
0.12359 | -64.1822
0.12391 | -64.2718
0.12422 | -64.3062
0.12453 | -64.3119
0.12484 | -64.2627
0.12515 | -64.1405
0.12546 | -64.0376
0.12577 | -63.9445
0.12608 | -63.8684
0.12639 | -63.7992
0.1267 | -63.7298
0.12701 | -63.6782
0.12732 | -63.6459
0.12763 | -63.6248
0.12794 | -63.6138
0.12825 | -63.6521
0.12856 | -63.7088
0.12887 | -63.7168
0.12918 | -63.7594
0.12949 | -63.8544
0.12981 | -63.9186
0.13012 | -63.9868
0.13043 | -64.0473
0.13074 | -64.0855
0.13105 | -64.018
0.13136 | -63.8882
0.13167 | -63.7262
0.13198 | -63.5286
0.13229 | -63.2892
0.1326 | -63.0468
0.13291 | -62.7746
0.13322 | -62.4936
0.13353 | -62.2984
0.13384 | -62.1213
0.13415 | -62.0639
0.13446 | -61.9525
0.13477 | -61.8928
0.13508 | -61.8458
0.13539 | -61.8482
0.1357 | -61.8736
0.13602 | -61.9188
0.13633 | -61.9891
0.13664 | -62.0837
0.13695 | -62.1835
0.13726 | -62.311
0.13757 | -62.4663
0.13788 | -62.6311
0.13819 | -62.8116
0.1385 | -63.0019
0.13881 | -63.2433
0.13912 | -63.4682
0.13943 | -63.6645
0.13974 | -63.8942
0.14005 | -64.1033
0.14036 | -64.2806
0.14067 | -64.4741
0.14098 | -64.6766
0.1413 | -64.8453
0.14161 | -64.9832
0.14192 | -65.1326
0.14223 | -65.2833
0.14254 | -65.4227
0.14285 | -65.6273
0.14316 | -65.7634
0.14347 | -65.9742
0.14378 | -66.1089
0.14409 | -66.27
0.1444 | -66.4251
0.14471 | -66.578
0.14502 | -66.667
0.14533 | -66.7379
0.14564 | -66.8232
0.14595 | -66.8684
0.14626 | -66.8891
0.14657 | -66.8928
0.14688 | -66.8732
0.14719 | -66.8265
0.1475 | -66.7633
0.14782 | -66.6442
0.14813 | -66.4977
0.14844 | -66.3731
0.14875 | -66.2511
0.14906 | -66.153
0.14937 | -66.0659
0.14968 | -65.9205
0.14999 | -65.8207
0.1503 | -65.7607
0.15061 | -65.7165
0.15092 | -65.6852
0.15123 | -65.6808
0.15154 | -65.6674
0.15185 | -65.6919
0.15216 | -65.7059
0.15247 | -65.7262
0.15279 | -65.7477
0.1531 | -65.7658
0.15341 | -65.7864
0.15372 | -65.8065
0.15403 | -65.8393
0.15434 | -65.8876
0.15465 | -65.9118
0.15496 | -65.9329
0.15527 | -65.9816
0.15558 | -66.0417
0.15589 | -66.1141
0.1562 | -66.2463
0.15651 | -66.3521
0.15682 | -66.4239
0.15713 | -66.5357
0.15744 | -66.6765
0.15775 | -66.8268
0.15806 | -66.988
0.15837 | -67.158
0.15868 | -67.3156
0.15899 | -67.431
0.15931 | -67.5566
0.15962 | -67.6141
0.15993 | -67.6646
0.16024 | -67.6519
0.16055 | -67.5724
0.16086 | -67.5546
0.16117 | -67.3622
0.16148 | -67.2424
0.16179 | -67.0945
0.1621 | -66.962
0.16241 | -66.8152
0.16272 | -66.6717
0.16303 | -66.5702
0.16334 | -66.4325
0.16365 | -66.3316
0.16396 | -66.2417
0.16427 | -66.1682
0.16459 | -66.1189
0.1649 | -66.0799
0.16521 | -66.0725
0.16552 | -66.0904
0.16583 | -66.1264
0.16614 | -66.1868
0.16645 | -66.2637
0.16676 | -66.3681
0.16707 | -66.4629
0.16738 | -66.5783
0.16769 | -66.6969
0.168 | -66.8012
0.16831 | -66.9159
0.16862 | -67.0586
0.16893 | -67.1655
0.16924 | -67.2536
0.16955 | -67.3435
0.16986 | -67.4087
0.17017 | -67.4277
0.17048 | -67.487
0.1708 | -67.5628
0.17111 | -67.6483
0.17142 | -67.6964
0.17173 | -67.7652
0.17204 | -67.8108
0.17235 | -67.8923
0.17266 | -67.9905
0.17297 | -68.0425
0.17328 | -68.1045
0.17359 | -68.1451
0.1739 | -68.18
0.17421 | -68.2382
0.17452 | -68.2964
0.17483 | -68.3619
0.17514 | -68.4276
0.17545 | -68.5026
0.17576 | -68.5848
0.17608 | -68.6487
0.17638 | -68.7077
0.1767 | -68.7568
0.17701 | -68.7987
0.17732 | -68.8555
0.17763 | -68.8866
0.17794 | -68.9132
0.17825 | -68.8833
0.17856 | -68.8104
0.17887 | -68.7609
0.17918 | -68.6672
0.17949 | -68.5964
0.1798 | -68.5164
0.18011 | -68.4189
0.18042 | -68.3303
0.18073 | -68.2447
0.18104 | -68.1402
0.18135 | -68.0348
0.18166 | -67.938
0.18197 | -67.8423
0.18229 | -67.7613
0.1826 | -67.6755
0.18291 | -67.5971
0.18322 | -67.5447
0.18353 | -67.4904
0.18384 | -67.4585
0.18415 | -67.4442
0.18446 | -67.4499
0.18477 | -67.464
0.18508 | -67.4942
0.18539 | -67.5461
0.1857 | -67.6121
0.18601 | -67.6961
0.18632 | -67.7923
0.18663 | -67.9079
0.18694 | -68.0408
0.18725 | -68.1558
0.18757 | -68.269
0.18787 | -68.3942
0.18819 | -68.5082
0.1885 | -68.6159
0.18881 | -68.7294
0.18912 | -68.8601
0.18943 | -68.997
0.18974 | -69.1245
0.19005 | -69.2511
0.19036 | -69.3935
0.19067 | -69.5408
0.19098 | -69.6771
0.19129 | -69.8108
0.1916 | -69.9958
0.19191 | -70.1682
0.19222 | -70.2715
0.19253 | -70.3735
0.19284 | -70.4581
0.19315 | -70.5654
0.19346 | -70.6808
0.19378 | -70.7582
0.19409 | -70.805
0.1944 | -70.8588
0.19471 | -70.8975
0.19502 | -70.916
0.19533 | -70.9164
0.19564 | -70.9094
0.19595 | -70.8827
0.19626 | -70.8822
0.19657 | -70.8802
0.19688 | -70.7999
0.19719 | -70.7465
0.1975 | -70.6962
0.19781 | -70.6413
0.19812 | -70.5912
0.19843 | -70.5477
0.19874 | -70.5288
0.19906 | -70.5238
0.19936 | -70.5081
0.19968 | -70.4988
0.19999 | -70.4897
0.2003 | -70.4954
0.20061 | -70.5222
0.20092 | -70.5687
0.20123 | -70.6087
0.20154 | -70.654
0.20185 | -70.7139
0.20216 | -70.7878
0.20247 | -70.8707
0.20278 | -70.9476
0.20309 | -71.0158
0.2034 | -71.0712
0.20371 | -71.1387
0.20402 | -71.2138
0.20433 | -71.279
0.20464 | -71.3152
0.20495 | -71.3495
0.20527 | -71.3629
0.20558 | -71.3755
0.20589 | -71.3475
0.2062 | -71.3604
0.20651 | -71.3854
0.20682 | -71.4083
0.20713 | -71.3913
0.20744 | -71.3822
0.20775 | -71.3542
0.20806 | -71.3383
0.20837 | -71.3444
0.20868 | -71.368
0.20899 | -71.3738
0.2093 | -71.3772
0.20961 | -71.4051
0.20992 | -71.4554
0.21023 | -71.5103
0.21055 | -71.5823
0.21085 | -71.6518
0.21117 | -71.7276
0.21148 | -71.8229
0.21179 | -71.8954
0.2121 | -71.9935
0.21241 | -72.0972
0.21272 | -72.2161
0.21303 | -72.2844
0.21334 | -72.3534
0.21365 | -72.4447
0.21396 | -72.5277
0.21427 | -72.565
0.21458 | -72.5899
0.21489 | -72.5208
0.2152 | -72.5305
0.21551 | -72.464
0.21582 | -72.3671
0.21613 | -72.2321
0.21644 | -72.1088
0.21676 | -71.992
0.21707 | -71.8474
0.21738 | -71.6941
0.21769 | -71.5417
0.218 | -71.3833
0.21831 | -71.2098
0.21862 | -71.0473
0.21893 | -70.8711
0.21924 | -70.7026
0.21955 | -70.5508
0.21986 | -70.4248
0.22017 | -70.2959
0.22048 | -70.1898
0.22079 | -70.0982
0.2211 | -70.0122
0.22141 | -69.9376
0.22172 | -69.8844
0.22204 | -69.8486
0.22234 | -69.838
0.22266 | -69.8435
0.22297 | -69.8554
0.22328 | -69.8542
0.22359 | -69.8716
0.2239 | -69.9737
0.22421 | -70.0284
0.22452 | -70.0904
0.22483 | -70.1641
0.22514 | -70.2479
0.22545 | -70.3339
0.22576 | -70.4187
0.22607 | -70.504
0.22638 | -70.586
0.22669 | -70.6703
0.227 | -70.7439
0.22731 | -70.8112
0.22762 | -70.8854
0.22793 | -70.9374
0.22825 | -70.9873
0.22856 | -71.0251
0.22887 | -71.0712
0.22918 | -71.1134
0.22949 | -71.1247
0.2298 | -71.1421
0.23011 | -71.1816
0.23042 | -71.2198
0.23073 | -71.2391
0.23104 | -71.2639
0.23135 | -71.3007
0.23166 | -71.3672
0.23197 | -71.439
0.23228 | -71.5017
0.23259 | -71.5455
0.2329 | -71.6385
0.23321 | -71.6841
0.23352 | -71.7403
0.23383 | -71.8044
0.23414 | -71.8709
0.23446 | -71.9473
0.23477 | -72.0347
0.23508 | -72.1179
0.23539 | -72.211
0.2357 | -72.2956
0.23601 | -72.3793
0.23632 | -72.4616
0.23663 | -72.5486
0.23694 | -72.6349
0.23725 | -72.7181
0.23756 | -72.7973
0.23787 | -72.8719
0.23818 | -72.9503
0.23849 | -73.0041
0.2388 | -73.0513
0.23911 | -73.0989
0.23942 | -73.156
0.23974 | -73.2661
0.24005 | -73.3166
0.24036 | -73.3174
0.24067 | -73.2985
0.24098 | -73.2552
0.24129 | -73.2414
0.2416 | -73.1989
0.24191 | -73.0887
0.24222 | -73.0272
0.24253 | -72.9576
0.24284 | -72.8887
0.24315 | -72.8073
0.24346 | -72.7185
0.24377 | -72.619
0.24408 | -72.5125
0.24439 | -72.4123
0.2447 | -72.3017
0.24501 | -72.1815
0.24532 | -72.0727
0.24563 | -71.955
0.24595 | -71.8589
0.24626 | -71.7572
0.24657 | -71.6682
0.24688 | -71.5829
0.24719 | -71.5019
0.2475 | -71.4387
0.24781 | -71.3824
0.24812 | -71.3438
0.24843 | -71.3092
0.24874 | -71.2937
0.24905 | -71.2841
0.24936 | -71.2786
0.24967 | -71.2794
0.24998 | -71.2894
0.25029 | -71.3204
0.2506 | -71.3997
0.25091 | -71.451
0.25122 | -71.5081
0.25154 | -71.564
0.25185 | -71.6322
0.25216 | -71.7027
0.25247 | -71.7661
0.25278 | -71.8246
0.25309 | -71.8913
0.2534 | -71.9528
0.25371 | -72.0179
0.25402 | -72.0792
0.25433 | -72.1422
0.25464 | -72.1939
0.25495 | -72.2292
0.25526 | -72.2737
0.25557 | -72.3231
0.25588 | -72.3577
0.25619 | -72.3972
0.2565 | -72.4347
0.25682 | -72.4533
0.25712 | -72.4774
0.25744 | -72.4946
0.25775 | -72.522
0.25806 | -72.5259
0.25837 | -72.5519
0.25868 | -72.5813
0.25899 | -72.5721
0.2593 | -72.5849
0.25961 | -72.607
0.25992 | -72.6376
0.26023 | -72.6459
0.26054 | -72.6547
0.26085 | -72.6695
0.26116 | -72.6856
0.26147 | -72.7071
0.26178 | -72.7078
0.26209 | -72.6964
0.2624 | -72.7057
0.26271 | -72.712
0.26303 | -72.7098
0.26334 | -72.7075
0.26365 | -72.6978
0.26396 | -72.7042
0.26427 | -72.7042
0.26458 | -72.6978
0.26489 | -72.69
0.2652 | -72.6732
0.26551 | -72.6589
0.26582 | -72.6446
0.26613 | -72.6331
0.26644 | -72.6149
0.26675 | -72.6044
0.26706 | -72.5953
0.26737 | -72.5891
0.26768 | -72.5948
0.26799 | -72.5961
0.26831 | -72.5959
0.26861 | -72.6023
0.26893 | -72.594
0.26924 | -72.5925
0.26955 | -72.598
0.26986 | -72.6033
0.27017 | -72.6141
0.27048 | -72.6201
0.27079 | -72.6241
0.2711 | -72.6381
0.27141 | -72.676
0.27172 | -72.7062
0.27203 | -72.7346
0.27234 | -72.7791
0.27265 | -72.8174
0.27296 | -72.8597
0.27327 | -72.9019
0.27358 | -72.9503
0.27389 | -73.0102
0.2742 | -73.0657
0.27452 | -73.1115
0.27483 | -73.1547
0.27514 | -73.1938
0.27545 | -73.223
0.27576 | -73.274
0.27607 | -73.3134
0.27638 | -73.3345
0.27669 | -73.3573
0.277 | -73.4007
0.27731 | -73.4283
0.27762 | -73.441
0.27793 | -73.4678
0.27824 | -73.5053
0.27855 | -73.5301
0.27886 | -73.5519
0.27917 | -73.572
0.27948 | -73.5847
0.2798 | -73.6091
0.2801 | -73.6466
0.28042 | -73.6748
0.28073 | -73.6931
0.28104 | -73.7178
0.28135 | -73.7704
0.28166 | -73.8446
0.28197 | -73.8706
0.28228 | -73.8909
0.28259 | -73.9325
0.2829 | -73.9794
0.28321 | -74.0275
0.28352 | -74.069
0.28383 | -74.1071
0.28414 | -74.1517
0.28445 | -74.2106
0.28476 | -74.2769
0.28507 | -74.3422
0.28538 | -74.4071
0.28569 | -74.4824
0.28601 | -74.5512
0.28632 | -74.6191
0.28663 | -74.6921
0.28694 | -74.7632
0.28725 | -74.8313
0.28756 | -74.9077
0.28787 | -74.9852
0.28818 | -75.0602
0.28849 | -75.1426
0.2888 | -75.2192
0.28911 | -75.2962
0.28942 | -75.3751
0.28973 | -75.4234
0.29004 | -75.4765
0.29035 | -75.5432
0.29066 | -75.6347
0.29097 | -75.7232
0.29128 | -75.797
0.29159 | -75.8467
0.29191 | -75.8978
0.29222 | -75.9267
0.29253 | -75.9614
0.29284 | -75.9975
0.29315 | -76.0335
0.29346 | -76.0672
0.29377 | -76.1005
0.29408 | -76.1321
0.29439 | -76.1496
0.2947 | -76.1699
0.29501 | -76.1877
0.29532 | -76.2224
0.29563 | -76.2176
0.29594 | -76.2468
0.29625 | -76.2228
0.29656 | -76.1885
0.29687 | -76.1911
0.29718 | -76.1924
0.2975 | -76.1766
0.29781 | -76.1441
0.29812 | -76.076
0.29843 | -76.0397
0.29874 | -75.9941
0.29905 | -75.9333
0.29936 | -75.8678
0.29967 | -75.8161
0.29998 | -75.7564
0.30029 | -75.6977
0.3006 | -75.6316
0.30091 | -75.5336
0.30122 | -75.4083
0.30153 | -75.3177
0.30184 | -75.218
0.30215 | -75.1117
0.30246 | -75.0128
0.30277 | -74.9033
0.30308 | -74.7779
0.3034 | -74.6536
0.30371 | -74.53
0.30402 | -74.411
0.30433 | -74.2719
0.30464 | -74.1601
0.30495 | -74.0542
0.30526 | -73.9335
0.30557 | -73.8033
0.30588 | -73.6901
0.30619 | -73.5867
0.3065 | -73.4844
0.30681 | -73.3885
0.30712 | -73.3019
0.30743 | -73.2183
0.30774 | -73.1418
0.30805 | -73.0683
0.30836 | -73.0026
0.30867 | -72.9457
0.30899 | -72.8969
0.3093 | -72.8579
0.30961 | -72.8183
0.30992 | -72.7938
0.31023 | -72.7727
0.31054 | -72.7573
0.31085 | -72.7537
0.31116 | -72.7674
0.31147 | -72.7878
0.31178 | -72.8041
0.31209 | -72.8106
0.3124 | -72.8289
0.31271 | -72.869
0.31302 | -72.9124
0.31333 | -73.0028
0.31364 | -73.0572
0.31395 | -73.1227
0.31426 | -73.1835
0.31457 | -73.253
0.31489 | -73.3256
0.3152 | -73.4041
0.31551 | -73.4811
0.31582 | -73.562
0.31613 | -73.6491
0.31644 | -73.7373
0.31675 | -73.8202
0.31706 | -73.8812
0.31737 | -73.9491
0.31768 | -74.022
0.31799 | -74.0838
0.3183 | -74.1617
0.31861 | -74.234
0.31892 | -74.2898
0.31923 | -74.3349
0.31954 | -74.3753
0.31985 | -74.4171
0.32016 | -74.4741
0.32048 | -74.51
0.32079 | -74.5325
0.3211 | -74.5544
0.32141 | -74.5686
0.32172 | -74.5623
0.32203 | -74.5539
0.32234 | -74.5323
0.32265 | -74.5293
0.32296 | -74.5109
0.32327 | -74.516
0.32358 | -74.5347
0.32389 | -74.512
0.3242 | -74.4954
0.32451 | -74.4815
0.32482 | -74.4625
0.32513 | -74.4553
0.32544 | -74.4569
0.32575 | -74.4464
0.32606 | -74.438
0.32638 | -74.4287
0.32669 | -74.4215
0.327 | -74.4146
0.32731 | -74.4112
0.32762 | -74.4159
0.32793 | -74.4356
0.32824 | -74.4612
0.32855 | -74.4743
0.32886 | -74.4827
0.32917 | -74.4908
0.32948 | -74.5001
0.32979 | -74.5165
0.3301 | -74.543
0.33041 | -74.5674
0.33072 | -74.5812
0.33103 | -74.6024
0.33134 | -74.6308
0.33165 | -74.6617
0.33197 | -74.6872
0.33227 | -74.7061
0.33259 | -74.7315
0.3329 | -74.7423
0.33321 | -74.7661
0.33352 | -74.7935
0.33383 | -74.8152
0.33414 | -74.8411
0.33445 | -74.8575
0.33476 | -74.8613
0.33507 | -74.8732
0.33538 | -74.8765
0.33569 | -74.8526
0.336 | -74.8456
0.33631 | -74.8367
0.33662 | -74.8272
0.33693 | -74.8247
0.33724 | -74.8158
0.33755 | -74.8045
0.33787 | -74.7864
0.33818 | -74.7602
0.33849 | -74.7382
0.3388 | -74.7201
0.33911 | -74.6962
0.33942 | -74.6755
0.33973 | -74.6444
0.34004 | -74.6247
0.34035 | -74.6064
0.34066 | -74.6165
0.34097 | -74.607
0.34128 | -74.5822
0.34159 | -74.5709
0.3419 | -74.5569
0.34221 | -74.5527
0.34252 | -74.5422
0.34283 | -74.5273
0.34314 | -74.5234
0.34346 | -74.5231
0.34376 | -74.5143
0.34408 | -74.5127
0.34439 | -74.5196
0.3447 | -74.5274
0.34501 | -74.5363
0.34532 | -74.5562
0.34563 | -74.5772
0.34594 | -74.5977
0.34625 | -74.6251
0.34656 | -74.6419
0.34687 | -74.6722
0.34718 | -74.6994
0.34749 | -74.7251
0.3478 | -74.7554
0.34811 | -74.792
0.34842 | -74.8257
0.34873 | -74.8529
0.34904 | -74.8881
0.34936 | -74.931
0.34967 | -74.9686
0.34998 | -75.0053
0.35029 | -75.0293
0.3506 | -75.0545
0.35091 | -75.0917
0.35122 | -75.1305
0.35153 | -75.1653
0.35184 | -75.1944
0.35215 | -75.2263
0.35246 | -75.2475
0.35277 | -75.2719
0.35308 | -75.3046
0.35339 | -75.3336
0.3537 | -75.3493
0.35401 | -75.3583
0.35432 | -75.3731
0.35463 | -75.3914
0.35495 | -75.417
0.35525 | -75.4319
0.35557 | -75.4227
0.35588 | -75.4172
0.35619 | -75.4106
0.3565 | -75.4038
0.35681 | -75.4041
0.35712 | -75.4037
0.35743 | -75.4078
0.35774 | -75.3985
0.35805 | -75.3875
0.35836 | -75.3859
0.35867 | -75.3802
0.35898 | -75.3739
0.35929 | -75.3714
0.3596 | -75.386
0.35991 | -75.4049
0.36022 | -75.4103
0.36053 | -75.4153
0.36085 | -75.4102
0.36116 | -75.4014
0.36147 | -75.3977
0.36178 | -75.3979
0.36209 | -75.4016
0.3624 | -75.4073
0.36271 | -75.4167
0.36302 | -75.431
0.36333 | -75.4413
0.36364 | -75.4644
0.36395 | -75.4911
0.36426 | -75.5075
0.36457 | -75.5397
0.36488 | -75.5699
0.36519 | -75.6
0.3655 | -75.6375
0.36581 | -75.6682
0.36612 | -75.71
0.36644 | -75.7542
0.36674 | -75.795
0.36706 | -75.8575
0.36737 | -75.9009
0.36768 | -75.9496
0.36799 | -75.992
0.3683 | -76.0477
0.36861 | -76.0832
0.36892 | -76.1343
0.36923 | -76.1984
0.36954 | -76.2524
0.36985 | -76.3124
0.37016 | -76.3657
0.37047 | -76.4235
0.37078 | -76.4789
0.37109 | -76.5312
0.3714 | -76.5776
0.37171 | -76.622
0.37202 | -76.6873
0.37234 | -76.737
0.37265 | -76.7695
0.37296 | -76.8073
0.37327 | -76.8635
0.37358 | -76.8907
0.37389 | -76.9169
0.3742 | -76.9266
0.37451 | -76.9299
0.37482 | -76.9248
0.37513 | -76.9252
0.37544 | -76.9298
0.37575 | -76.9289
0.37606 | -76.8882
0.37637 | -76.8895
0.37668 | -76.8761
0.37699 | -76.8574
0.3773 | -76.8233
0.37761 | -76.7889
0.37792 | -76.7392
0.37823 | -76.695
0.37855 | -76.6692
0.37886 | -76.6258
0.37917 | -76.5824
0.37948 | -76.545
0.37979 | -76.5164
0.3801 | -76.4795
0.38041 | -76.4297
0.38072 | -76.3816
0.38103 | -76.3304
0.38134 | -76.2878
0.38165 | -76.2584
0.38196 | -76.2355
0.38227 | -76.1963
0.38258 | -76.1672
0.38289 | -76.1343
0.3832 | -76.0999
0.38351 | -76.067
0.38383 | -76.0442
0.38414 | -76.0227
0.38445 | -75.999
0.38476 | -75.9739
0.38507 | -75.9656
0.38538 | -75.9526
0.38569 | -75.9423
0.386 | -75.9355
0.38631 | -75.918
0.38662 | -75.9108
0.38693 | -75.9168
0.38724 | -75.9237
0.38755 | -75.9224
0.38786 | -75.9317
0.38817 | -75.9402
0.38848 | -75.9406
0.38879 | -75.9476
0.3891 | -75.963
0.38941 | -75.9739
0.38972 | -75.9925
0.39004 | -76.0167
0.39035 | -76.0293
0.39066 | -76.0492
0.39097 | -76.0733
0.39128 | -76.0961
0.39159 | -76.1195
0.3919 | -76.1432
0.39221 | -76.1667
0.39252 | -76.1917
0.39283 | -76.2188
0.39314 | -76.2472
0.39345 | -76.2708
0.39376 | -76.2893
0.39407 | -76.3065
0.39438 | -76.3293
0.39469 | -76.3486
0.395 | -76.368
0.39532 | -76.3922
0.39563 | -76.4185
0.39594 | -76.4562
0.39625 | -76.4732
0.39656 | -76.4718
0.39687 | -76.483
0.39718 | -76.497
0.39749 | -76.5159
0.3978 | -76.5284
0.39811 | -76.5412
0.39842 | -76.5568
0.39873 | -76.5577
0.39904 | -76.5706
0.39935 | -76.5758
0.39966 | -76.5799
0.39997 | -76.5981
0.40028 | -76.6095
0.40059 | -76.6126
0.4009 | -76.6203
0.40121 | -76.6251
0.40153 | -76.6456
0.40184 | -76.6577
0.40215 | -76.6699
0.40246 | -76.6765
0.40277 | -76.6793
0.40308 | -76.6747
0.40339 | -76.6768
0.4037 | -76.6788
0.40401 | -76.6859
0.40432 | -76.6883
0.40463 | -76.6923
0.40494 | -76.6928
0.40525 | -76.7054
0.40556 | -76.7183
0.40587 | -76.7409
0.40618 | -76.7492
0.40649 | -76.755
0.40681 | -76.7641
0.40712 | -76.7699
0.40743 | -76.7783
0.40774 | -76.7887
0.40805 | -76.8054
0.40836 | -76.8224
0.40867 | -76.8392
0.40898 | -76.8633
0.40929 | -76.8804
0.4096 | -76.9067
0.40991 | -76.9331
0.41022 | -76.9604
0.41053 | -76.9717
0.41084 | -77.0005
0.41115 | -77.0305
0.41146 | -77.0686
0.41177 | -77.0973
0.41208 | -77.139
0.41239 | -77.1705
0.4127 | -77.2112
0.41302 | -77.2575
0.41333 | -77.3116
0.41364 | -77.3466
0.41395 | -77.3834
0.41426 | -77.4245
0.41457 | -77.4636
0.41488 | -77.5064
0.41519 | -77.5453
0.4155 | -77.58
0.41581 | -77.6105
0.41612 | -77.6563
0.41643 | -77.6929
0.41674 | -77.7282
0.41705 | -77.7759
0.41736 | -77.8167
0.41767 | -77.8423
0.41798 | -77.8778
0.41829 | -77.9183
0.41861 | -77.9459
0.41891 | -77.9784
0.41923 | -77.9755
0.41954 | -77.9799
0.41985 | -78.0032
0.42016 | -78.0228
0.42047 | -78.037
0.42078 | -78.0509
0.42109 | -78.0422
0.4214 | -78.0549
0.42171 | -78.0603
0.42202 | -78.0643
0.42233 | -78.0732
0.42264 | -78.0834
0.42295 | -78.0961
0.42326 | -78.0859
0.42357 | -78.089
0.42388 | -78.0878
0.42419 | -78.0852
0.42451 | -78.0717
0.42482 | -78.0563
0.42513 | -78.0428
0.42544 | -78.0368
0.42575 | -78.0493
0.42606 | -78.0338
0.42637 | -77.9957
0.42668 | -77.9684
0.42699 | -77.9453
0.4273 | -77.921
0.42761 | -77.9046
0.42792 | -77.8959
0.42823 | -77.8732
0.42854 | -77.8512
0.42885 | -77.8379
0.42916 | -77.7996
0.42947 | -77.7605
0.42978 | -77.7428
0.4301 | -77.7218
0.4304 | -77.7014
0.43072 | -77.683
0.43103 | -77.667
0.43134 | -77.6541
0.43165 | -77.6421
0.43196 | -77.6264
0.43227 | -77.6096
0.43258 | -77.6042
0.43289 | -77.6108
0.4332 | -77.6084
0.43351 | -77.6016
0.43382 | -77.5937
0.43413 | -77.5916
0.43444 | -77.5935
0.43475 | -77.5961
0.43506 | -77.5976
0.43537 | -77.5956
0.43568 | -77.6012
0.436 | -77.6144
0.43631 | -77.6227
0.43662 | -77.6318
0.43693 | -77.64
0.43724 | -77.6536
0.43755 | -77.6625
0.43786 | -77.6704
0.43817 | -77.6851
0.43848 | -77.7051
0.43879 | -77.7415
0.4391 | -77.7618
0.43941 | -77.7752
0.43972 | -77.7967
0.44003 | -77.8121
0.44034 | -77.8321
0.44065 | -77.8736
0.44096 | -77.8906
0.44127 | -77.9045
0.44159 | -77.9261
0.44189 | -77.9532
0.44221 | -77.9798
0.44252 | -78.0039
0.44283 | -78.0283
0.44314 | -78.0479
0.44345 | -78.0582
0.44376 | -78.0706
0.44407 | -78.095
0.44438 | -78.1092
0.44469 | -78.1141
0.445 | -78.1279
0.44531 | -78.1576
0.44562 | -78.1801
0.44593 | -78.1933
0.44624 | -78.2082
0.44655 | -78.216
0.44686 | -78.2084
0.44717 | -78.2107
0.44749 | -78.2135
0.44779 | -78.1892
0.44811 | -78.1862
0.44842 | -78.1937
0.44873 | -78.1867
0.44904 | -78.1808
0.44935 | -78.1773
0.44966 | -78.1712
0.44997 | -78.1914
0.45028 | -78.1802
0.45059 | -78.1453
0.4509 | -78.1285
0.45121 | -78.1175
0.45152 | -78.1075
0.45183 | -78.0984
0.45214 | -78.0823
0.45245 | -78.0645
0.45276 | -78.0513
0.45308 | -78.0378
0.45338 | -78.0262
0.4537 | -78.012
0.45401 | -77.9974
0.45432 | -77.989
0.45463 | -77.975
0.45494 | -77.9621
0.45525 | -77.9474
0.45556 | -77.9428
0.45587 | -77.9301
0.45618 | -77.9158
0.45649 | -77.8983
0.4568 | -77.8987
0.45711 | -77.8955
0.45742 | -77.8978
0.45773 | -77.9064
0.45804 | -77.908
0.45835 | -77.9077
0.45866 | -77.9097
0.45898 | -77.8955
0.45928 | -77.8997
0.4596 | -77.905
0.45991 | -77.9113
0.46022 | -77.9169
0.46053 | -77.933
0.46084 | -77.955
0.46115 | -77.9654
0.46146 | -77.9841
0.46177 | -77.9954
0.46208 | -78.0118
0.46239 | -78.0273
0.4627 | -78.0424
0.46301 | -78.0652
0.46332 | -78.0864
0.46363 | -78.1127
0.46394 | -78.1438
0.46425 | -78.1701
0.46457 | -78.1941
0.46487 | -78.2163
0.46519 | -78.2499
0.4655 | -78.2798
0.46581 | -78.3172
0.46612 | -78.3524
0.46643 | -78.3816
0.46674 | -78.4045
0.46705 | -78.4374
0.46736 | -78.4764
0.46767 | -78.511
0.46798 | -78.5434
0.46829 | -78.5802
0.4686 | -78.6166
0.46891 | -78.6513
0.46922 | -78.6574
0.46953 | -78.6718
0.46984 | -78.7033
0.47015 | -78.7182
0.47047 | -78.7292
0.47077 | -78.7438
0.47109 | -78.7471
0.4714 | -78.7645
0.47171 | -78.785
0.47202 | -78.7791
0.47233 | -78.7861
0.47264 | -78.7984
0.47295 | -78.7986
0.47326 | -78.8018
0.47357 | -78.8041
0.47388 | -78.7985
0.47419 | -78.7818
0.4745 | -78.768
0.47481 | -78.7394
0.47512 | -78.7282
0.47543 | -78.7128
0.47574 | -78.6969
0.47605 | -78.6897
0.47636 | -78.6822
0.47668 | -78.6593
0.47699 | -78.6408
0.4773 | -78.6288
0.47761 | -78.6029
0.47792 | -78.5735
0.47823 | -78.5482
0.47854 | -78.523
0.47885 | -78.501
0.47916 | -78.4731
0.47947 | -78.4634
0.47978 | -78.4436
0.48009 | -78.4161
0.4804 | -78.4007
0.48071 | -78.3824
0.48102 | -78.3589
0.48133 | -78.3377
0.48164 | -78.3198
0.48196 | -78.3006
0.48226 | -78.2812
0.48258 | -78.2587
0.48289 | -78.2374
0.4832 | -78.2403
0.48351 | -78.2466
0.48382 | -78.2621
0.48413 | -78.2496
0.48444 | -78.2401
0.48475 | -78.2343
0.48506 | -78.2328
0.48537 | -78.2337
0.48568 | -78.2401
0.48599 | -78.2463
0.4863 | -78.2552
0.48661 | -78.2661
0.48692 | -78.2755
0.48723 | -78.2842
0.48754 | -78.3013
0.48785 | -78.322
0.48817 | -78.3433
0.48848 | -78.3671
0.48879 | -78.3962
0.4891 | -78.4244
0.48941 | -78.4563
0.48972 | -78.4851
0.49003 | -78.5091
0.49034 | -78.5376
0.49065 | -78.5626
0.49096 | -78.5944
0.49127 | -78.623
0.49158 | -78.6548
0.49189 | -78.6868
0.4922 | -78.7197
0.49251 | -78.7408
0.49282 | -78.7862
0.49313 | -78.8183
0.49345 | -78.8469
0.49375 | -78.87
0.49407 | -78.9039
0.49438 | -78.9433
0.49469 | -78.9725
0.495 | -78.9944
0.49531 | -79.0176
0.49562 | -79.0399
0.49593 | -79.0506
0.49624 | -79.07
0.49655 | -79.0912
0.49686 | -79.0983
0.49717 | -79.1144
0.49748 | -79.1264
0.49779 | -79.1394
0.4981 | -79.1488
0.49841 | -79.1655
0.49872 | -79.1809
0.49903 | -79.1865
0.49934 | -79.1999
0.49966 | -79.203
0.49997 | -79.2056
0.50028 | -79.2053
0.50059 | -79.2106
0.5009 | -79.2226
0.50121 | -79.218
0.50152 | -79.2037
0.50183 | -79.1941
0.50214 | -79.1853
0.50245 | -79.1799
0.50276 | -79.1716
0.50307 | -79.1679
0.50338 | -79.1636
0.50369 | -79.1779
0.504 | -79.186
0.50431 | -79.1743
0.50462 | -79.1716
0.50493 | -79.1645
0.50525 | -79.1552
0.50556 | -79.1486
0.50587 | -79.1309
0.50618 | -79.1215
0.50649 | -79.1249
0.5068 | -79.1111
0.50711 | -79.0966
0.50742 | -79.0825
0.50773 | -79.0766
0.50804 | -79.0742
0.50835 | -79.0701
0.50866 | -79.0749
0.50897 | -79.0738
0.50928 | -79.0748
0.50959 | -79.075
0.5099 | -79.0749
0.51021 | -79.082
0.51052 | -79.1028
0.51084 | -79.1095
0.51114 | -79.1169
0.51146 | -79.1259
0.51177 | -79.1371
0.51208 | -79.1471
0.51239 | -79.1606
0.5127 | -79.1771
0.51301 | -79.1971
0.51332 | -79.2081
0.51363 | -79.2235
0.51394 | -79.2415
0.51425 | -79.2547
0.51456 | -79.2641
0.51487 | -79.28
0.51518 | -79.3084
0.51549 | -79.3273
0.5158 | -79.3372
0.51611 | -79.3631
0.51642 | -79.3843
0.51674 | -79.4129
0.51704 | -79.4379
0.51736 | -79.4649
0.51767 | -79.4982
0.51798 | -79.5245
0.51829 | -79.5458
0.5186 | -79.5659
0.51891 | -79.588
0.51922 | -79.6099
0.51953 | -79.6313
0.51984 | -79.6583
0.52015 | -79.6809
0.52046 | -79.7032
0.52077 | -79.7216
0.52108 | -79.7372
0.52139 | -79.7535
0.5217 | -79.7692
0.52201 | -79.7871
0.52233 | -79.8009
0.52263 | -79.8159
0.52295 | -79.842
0.52326 | -79.8572
0.52357 | -79.8688
0.52388 | -79.8836
0.52419 | -79.8968
0.5245 | -79.9088
0.52481 | -79.9212
0.52512 | -79.9274
0.52543 | -79.9447
0.52574 | -79.9661
0.52605 | -79.9797
0.52636 | -79.9839
0.52667 | -79.9886
0.52698 | -79.9984
0.52729 | -79.9993
0.5276 | -80.0076
0.52791 | -80.0156
0.52822 | -80.0216
0.52853 | -79.9747
0.52885 | -79.977
0.52916 | -79.9686
0.52947 | -79.9589
0.52978 | -79.9567
0.53009 | -79.9593
0.5304 | -79.9521
0.53071 | -79.9446
0.53102 | -79.9381
0.53133 | -79.9369
0.53164 | -79.9312
0.53195 | -79.931
0.53226 | -79.9391
0.53257 | -79.9338
0.53288 | -79.9284
0.53319 | -79.9288
0.5335 | -79.9219
0.53381 | -79.9127
0.53412 | -79.8937
0.53444 | -79.8749
0.53475 | -79.858
0.53506 | -79.8388
0.53537 | -79.8217
0.53568 | -79.7976
0.53599 | -79.7641
0.5363 | -79.7413
0.53661 | -79.7208
0.53692 | -79.6965
0.53723 | -79.6728
0.53754 | -79.6545
0.53785 | -79.6246
0.53816 | -79.6003
0.53847 | -79.5835
0.53878 | -79.5623
0.53909 | -79.5441
0.5394 | -79.5276
0.53971 | -79.5055
0.54002 | -79.4808
0.54034 | -79.4529
0.54065 | -79.4292
0.54096 | -79.4065
0.54127 | -79.3705
0.54158 | -79.3374
0.54189 | -79.3115
0.5422 | -79.2856
0.54251 | -79.2703
0.54282 | -79.2535
0.54313 | -79.2297
0.54344 | -79.204
0.54375 | -79.1878
0.54406 | -79.1658
0.54437 | -79.1426
0.54468 | -79.1173
0.54499 | -79.0885
0.5453 | -79.0607
0.54561 | -79.0377
0.54592 | -79.0156
0.54624 | -79.0017
0.54655 | -78.9785
0.54686 | -78.9582
0.54717 | -78.9399
0.54748 | -78.9207
0.54779 | -78.8974
0.5481 | -78.8731
0.54841 | -78.8578
0.54872 | -78.8413
0.54903 | -78.8242
0.54934 | -78.8031
0.54965 | -78.7799
0.54996 | -78.7651
0.55027 | -78.7535
0.55058 | -78.7363
0.55089 | -78.7291
0.5512 | -78.7159
0.55151 | -78.7006
0.55183 | -78.6988
0.55214 | -78.6947
0.55245 | -78.6877
0.55276 | -78.678
0.55307 | -78.676
0.55338 | -78.6742
0.55369 | -78.6586
0.554 | -78.6518
0.55431 | -78.6439
0.55462 | -78.6407
0.55493 | -78.6343
0.55524 | -78.6438
0.55555 | -78.6391
0.55586 | -78.6415
0.55617 | -78.6436
0.55648 | -78.6467
0.55679 | -78.6505
0.5571 | -78.6554
0.55741 | -78.6532
0.55773 | -78.6561
0.55804 | -78.6637
0.55835 | -78.671
0.55866 | -78.6817
0.55897 | -78.6936
0.55928 | -78.7096
0.55959 | -78.7158
0.5599 | -78.7253
0.56021 | -78.7342
0.56052 | -78.7379
0.56083 | -78.7579
0.56114 | -78.7641
0.56145 | -78.7769
0.56176 | -78.7897
0.56207 | -78.8071
0.56238 | -78.8219
0.56269 | -78.8332
0.563 | -78.8524
0.56332 | -78.8713
0.56363 | -78.8912
0.56394 | -78.9046
0.56425 | -78.923
0.56456 | -78.9379
0.56487 | -78.9473
0.56518 | -78.9567
0.56549 | -78.9702
0.5658 | -78.9837
0.56611 | -78.9977
0.56642 | -79.0118
0.56673 | -79.0292
0.56704 | -79.0643
0.56735 | -79.0879
0.56766 | -79.1022
0.56797 | -79.119
0.56828 | -79.1325
0.56859 | -79.1413
0.5689 | -79.1506
0.56922 | -79.1621
0.56953 | -79.1727
0.56984 | -79.1749
0.57015 | -79.1834
0.57046 | -79.1935
0.57077 | -79.2044
0.57108 | -79.2153
0.57139 | -79.2265
0.5717 | -79.2358
0.57201 | -79.2466
0.57232 | -79.2537
0.57263 | -79.2645
0.57294 | -79.2831
0.57325 | -79.2991
0.57356 | -79.3181
0.57387 | -79.3328
0.57418 | -79.3511
0.57449 | -79.3672
0.57481 | -79.3903
0.57512 | -79.4183
0.57543 | -79.4351
0.57574 | -79.4484
0.57605 | -79.4697
0.57636 | -79.4873
0.57667 | -79.4986
0.57698 | -79.5144
0.57729 | -79.5393
0.5776 | -79.5629
0.57791 | -79.5818
0.57822 | -79.6004
0.57853 | -79.6205
0.57884 | -79.6402
0.57915 | -79.6644
0.57946 | -79.6875
0.57977 | -79.7157
0.58008 | -79.7433
0.58039 | -79.7689
0.58071 | -79.8095
0.58102 | -79.8466
0.58133 | -79.8808
0.58164 | -79.9122
0.58195 | -79.9401
0.58226 | -79.9693
0.58257 | -80.0002
0.58288 | -80.0333
0.58319 | -80.0678
0.5835 | -80.0993
0.58381 | -80.1352
0.58412 | -80.1736
0.58443 | -80.2094
0.58474 | -80.2509
0.58505 | -80.2897
0.58536 | -80.3261
0.58567 | -80.3655
0.58598 | -80.404
0.5863 | -80.4431
0.58661 | -80.4871
0.58692 | -80.5288
0.58723 | -80.5717
0.58754 | -80.615
0.58785 | -80.6621
0.58816 | -80.7117
0.58847 | -80.7505
0.58878 | -80.7893
0.58909 | -80.8295
0.5894 | -80.8825
0.58971 | -80.9397
0.59002 | -80.9796
0.59033 | -81.0223
0.59064 | -81.0693
0.59095 | -81.1063
0.59126 | -81.1421
0.59157 | -81.1822
0.59188 | -81.2282
0.5922 | -81.2667
0.59251 | -81.3108
0.59282 | -81.3504
0.59313 | -81.3859
0.59344 | -81.4193
0.59375 | -81.456
0.59406 | -81.5014
0.59437 | -81.5298
0.59468 | -81.5673
0.59499 | -81.5933
0.5953 | -81.6318
0.59561 | -81.6646
0.59592 | -81.695
0.59623 | -81.7334
0.59654 | -81.7577
0.59685 | -81.7865
0.59716 | -81.8078
0.59747 | -81.8322
0.59779 | -81.8548
0.5981 | -81.8945
0.59841 | -81.9225
0.59872 | -81.9382
0.59903 | -81.9635
0.59934 | -81.981
0.59965 | -81.9928
0.59996 | -82.0042
0.60027 | -82.0062
0.60058 | -82.0376
0.60089 | -82.062
0.6012 | -82.0814
0.60151 | -82.1158
0.60182 | -82.1209
0.60213 | -82.135
0.60244 | -82.1462
0.60275 | -82.1596
0.60306 | -82.1662
0.60337 | -82.1711
0.60369 | -82.1838
0.604 | -82.1849
0.60431 | -82.19
0.60462 | -82.1915
0.60493 | -82.1933
0.60524 | -82.1912
0.60555 | -82.1896
0.60586 | -82.1897
0.60617 | -82.2158
0.60648 | -82.2244
0.60679 | -82.225
0.6071 | -82.2297
0.60741 | -82.2287
0.60772 | -82.2277
0.60803 | -82.2274
0.60834 | -82.2296
0.60865 | -82.2283
0.60896 | -82.2277
0.60928 | -82.215
0.60959 | -82.2157
0.6099 | -82.2165
0.61021 | -82.2206
0.61052 | -82.2182
0.61083 | -82.2179
0.61114 | -82.2184
0.61145 | -82.2336
0.61176 | -82.24
0.61207 | -82.2438
0.61238 | -82.2417
0.61269 | -82.2377
0.613 | -82.2418
0.61331 | -82.2376
0.61362 | -82.2356
0.61393 | -82.2448
0.61424 | -82.2479
0.61455 | -82.2495
0.61486 | -82.2507
0.61518 | -82.2542
0.61549 | -82.2563
0.6158 | -82.2597
0.61611 | -82.2602
0.61642 | -82.2586
0.61673 | -82.2622
0.61704 | -82.2702
0.61735 | -82.2693
0.61766 | -82.2768
0.61797 | -82.2873
0.61828 | -82.2883
0.61859 | -82.2928
0.6189 | -82.3029
0.61921 | -82.31
0.61952 | -82.3185
0.61983 | -82.3095
0.62014 | -82.3067
0.62045 | -82.318
0.62077 | -82.3188
0.62108 | -82.3157
0.62139 | -82.3128
0.6217 | -82.3055
0.62201 | -82.3082
0.62232 | -82.31
0.62263 | -82.3132
0.62294 | -82.3124
0.62325 | -82.3125
0.62356 | -82.2878
0.62387 | -82.2833
0.62418 | -82.2681
0.62449 | -82.2755
0.6248 | -82.2726
0.62511 | -82.2634
0.62542 | -82.2534
0.62573 | -82.2474
0.62604 | -82.2402
0.62635 | -82.2132
0.62667 | -82.201
0.62698 | -82.1877
0.62729 | -82.1648
0.6276 | -82.1403
0.62791 | -82.1268
0.62822 | -82.1095
0.62853 | -82.0942
0.62884 | -82.0805
0.62915 | -82.0655
0.62946 | -82.0468
0.62977 | -82.0298
0.63008 | -82.0105
0.63039 | -81.99
0.6307 | -81.9773
0.63101 | -81.9595
0.63132 | -81.9448
0.63163 | -81.9221
0.63194 | -81.9037
0.63226 | -81.8843
0.63257 | -81.857
0.63288 | -81.8364
0.63319 | -81.8065
0.6335 | -81.7844
0.63381 | -81.7585
0.63412 | -81.7338
0.63443 | -81.7116
0.63474 | -81.6898
0.63505 | -81.669
0.63536 | -81.6406
0.63567 | -81.6096
0.63598 | -81.5534
0.63629 | -81.4947
0.6366 | -81.463
0.63691 | -81.4286
0.63722 | -81.3908
0.63753 | -81.3589
0.63784 | -81.3241
0.63816 | -81.2923
0.63847 | -81.2638
0.63878 | -81.2243
0.63909 | -81.1858
0.6394 | -81.1509
0.63971 | -81.1195
0.64002 | -81.0878
0.64033 | -81.0523
0.64064 | -81.0219
0.64095 | -80.9806
0.64126 | -80.9404
0.64157 | -80.907
0.64188 | -80.8763
0.64219 | -80.8456
0.6425 | -80.813
0.64281 | -80.7819
0.64312 | -80.7571
0.64343 | -80.7305
0.64375 | -80.7026
0.64406 | -80.6725
0.64437 | -80.6411
0.64468 | -80.6112
0.64499 | -80.5881
0.6453 | -80.56
0.64561 | -80.5334
0.64592 | -80.503
0.64623 | -80.475
0.64654 | -80.4479
0.64685 | -80.4237
0.64716 | -80.4001
0.64747 | -80.3782
0.64778 | -80.3532
0.64809 | -80.3276
0.6484 | -80.307
0.64871 | -80.2881
0.64902 | -80.2729
0.64933 | -80.2586
0.64965 | -80.2432
0.64996 | -80.2256
0.65027 | -80.2041
0.65058 | -80.1852
0.65089 | -80.1724
0.6512 | -80.1588
0.65151 | -80.1449
0.65182 | -80.1311
0.65213 | -80.1172
0.65244 | -80.1055
0.65275 | -80.0966
0.65306 | -80.0898
0.65337 | -80.0867
0.65368 | -80.0855
0.65399 | -80.0876
0.6543 | -80.0789
0.65461 | -80.0749
0.65492 | -80.0726
0.65524 | -80.074
0.65555 | -80.0739
0.65586 | -80.0735
0.65617 | -80.0752
0.65648 | -80.0777
0.65679 | -80.0812
0.6571 | -80.089
0.65741 | -80.0973
0.65772 | -80.1028
0.65803 | -80.1181
0.65834 | -80.1296
0.65865 | -80.1417
0.65896 | -80.1564
0.65927 | -80.1697
0.65958 | -80.1835
0.65989 | -80.2004
0.6602 | -80.2193
0.66051 | -80.2382
0.66082 | -80.2601
0.66114 | -80.2828
0.66144 | -80.3056
0.66176 | -80.3295
0.66207 | -80.3566
0.66238 | -80.3831
0.66269 | -80.4098
0.663 | -80.4522
0.66331 | -80.4773
0.66362 | -80.5025
0.66393 | -80.5309
0.66424 | -80.5622
0.66455 | -80.594
0.66486 | -80.6264
0.66517 | -80.6657
0.66548 | -80.6995
0.66579 | -80.7342
0.6661 | -80.7665
0.66641 | -80.8033
0.66673 | -80.8346
0.66704 | -80.8734
0.66735 | -80.9095
0.66766 | -80.9442
0.66797 | -80.9746
0.66828 | -81.0068
0.66859 | -81.0436
0.6689 | -81.0973
0.66921 | -81.1323
0.66952 | -81.1774
0.66983 | -81.2159
0.67014 | -81.2515
0.67045 | -81.2857
0.67076 | -81.3207
0.67107 | -81.3536
0.67138 | -81.392
0.67169 | -81.4113
0.672 | -81.4569
0.67231 | -81.4802
0.67263 | -81.5091
0.67293 | -81.5372
0.67325 | -81.5688
0.67356 | -81.5874
0.67387 | -81.5991
0.67418 | -81.6225
0.67449 | -81.6419
0.6748 | -81.6607
0.67511 | -81.6807
0.67542 | -81.7021
0.67573 | -81.7207
0.67604 | -81.7407
0.67635 | -81.7579
0.67666 | -81.7688
0.67697 | -81.7854
0.67728 | -81.795
0.67759 | -81.8096
0.6779 | -81.8167
0.67822 | -81.8233
0.67853 | -81.8297
0.67884 | -81.8327
0.67915 | -81.8414
0.67946 | -81.8437
0.67977 | -81.8504
0.68008 | -81.848
0.68039 | -81.8468
0.6807 | -81.8282
0.68101 | -81.8301
0.68132 | -81.8265
0.68163 | -81.826
0.68194 | -81.8247
0.68225 | -81.8301
0.68256 | -81.8399
0.68287 | -81.8619
0.68318 | -81.8542
0.68349 | -81.8493
0.6838 | -81.8462
0.68412 | -81.8356
0.68442 | -81.8292
0.68474 | -81.8152
0.68505 | -81.7925
0.68536 | -81.7721
0.68567 | -81.7628
0.68598 | -81.7471
0.68629 | -81.7262
0.6866 | -81.708
0.68691 | -81.6926
0.68722 | -81.6777
0.68753 | -81.6717
0.68784 | -81.6662
0.68815 | -81.6485
0.68846 | -81.6367
0.68877 | -81.6244
0.68908 | -81.6106
0.68939 | -81.5919
0.6897 | -81.5868
0.69002 | -81.565
0.69033 | -81.5537
0.69064 | -81.5419
0.69095 | -81.5351
0.69126 | -81.5264
0.69157 | -81.5163
0.69188 | -81.5125
0.69219 | -81.5052
0.6925 | -81.491
0.69281 | -81.4691
0.69312 | -81.4674
0.69343 | -81.4677
0.69374 | -81.4655
0.69405 | -81.4577
0.69436 | -81.4532
0.69467 | -81.4484
0.69498 | -81.4463
0.69529 | -81.4527
0.69561 | -81.4501
0.69591 | -81.4512
0.69623 | -81.4521
0.69654 | -81.4552
0.69685 | -81.4588
0.69716 | -81.4609
0.69747 | -81.4658
0.69778 | -81.4721
0.69809 | -81.4733
0.6984 | -81.4757
0.69871 | -81.4913
0.69902 | -81.4981
0.69933 | -81.5041
0.69964 | -81.5111
0.69995 | -81.5181
0.70026 | -81.524
0.70057 | -81.5251
0.70088 | -81.5326
0.70119 | -81.5458
0.70151 | -81.5546
0.70181 | -81.5636
0.70213 | -81.5672
0.70244 | -81.5725
0.70275 | -81.5775
0.70306 | -81.5843
0.70337 | -81.5926
0.70368 | -81.6028
0.70399 | -81.6124
0.7043 | -81.6125
0.70461 | -81.6199
0.70492 | -81.622
0.70523 | -81.6287
0.70554 | -81.6374
0.70585 | -81.636
0.70616 | -81.639
0.70647 | -81.6444
0.70678 | -81.6511
0.7071 | -81.6542
0.7074 | -81.6547
0.70772 | -81.6546
0.70803 | -81.6584
0.70834 | -81.6603
0.70865 | -81.6641
0.70896 | -81.668
0.70927 | -81.6721
0.70958 | -81.6727
0.70989 | -81.673
0.7102 | -81.678
0.71051 | -81.6805
0.71082 | -81.6808
0.71113 | -81.6796
0.71144 | -81.6812
0.71175 | -81.6805
0.71206 | -81.6897
0.71237 | -81.6874
0.71268 | -81.6871
0.713 | -81.6877
0.7133 | -81.6827
0.71362 | -81.6858
0.71393 | -81.6884
0.71424 | -81.6885
0.71455 | -81.6935
0.71486 | -81.6951
0.71517 | -81.7034
0.71548 | -81.7081
0.71579 | -81.7042
0.7161 | -81.6984
0.71641 | -81.7042
0.71672 | -81.7014
0.71703 | -81.7054
0.71734 | -81.7025
0.71765 | -81.6981
0.71796 | -81.6986
0.71827 | -81.6985
0.71858 | -81.6948
0.71889 | -81.6985
0.71921 | -81.7014
0.71952 | -81.7027
0.71983 | -81.7072
0.72014 | -81.7106
0.72045 | -81.717
0.72076 | -81.7249
0.72107 | -81.726
0.72138 | -81.7341
0.72169 | -81.7376
0.722 | -81.7427
0.72231 | -81.7527
0.72262 | -81.7742
0.72293 | -81.7863
0.72324 | -81.7894
0.72355 | -81.7938
0.72386 | -81.7995
0.72417 | -81.8033
0.72449 | -81.8058
0.72479 | -81.8116
0.72511 | -81.817
0.72542 | -81.8237
0.72573 | -81.8411
0.72604 | -81.8519
0.72635 | -81.8616
0.72666 | -81.8733
0.72697 | -81.8845
0.72728 | -81.8973
0.72759 | -81.9079
0.7279 | -81.9219
0.72821 | -81.9334
0.72852 | -81.9439
0.72883 | -81.9513
0.72914 | -81.9642
0.72945 | -81.9798
0.72976 | -81.9914
0.73007 | -82.0066
0.73038 | -82.0228
0.73069 | -82.0357
0.73101 | -82.0495
0.73132 | -82.0645
0.73163 | -82.0817
0.73194 | -82.0984
0.73225 | -82.1126
0.73256 | -82.1296
0.73287 | -82.149
0.73318 | -82.1697
0.73349 | -82.1886
0.7338 | -82.2087
0.73411 | -82.2258
0.73442 | -82.2454
0.73473 | -82.2646
0.73504 | -82.2916
0.73535 | -82.3239
0.73566 | -82.3454
0.73598 | -82.3664
0.73628 | -82.3849
0.7366 | -82.4079
0.73691 | -82.4293
0.73722 | -82.4424
0.73753 | -82.4633
0.73784 | -82.4853
0.73815 | -82.511
0.73846 | -82.5307
0.73877 | -82.5498
0.73908 | -82.5732
0.73939 | -82.5997
0.7397 | -82.625
0.74001 | -82.6533
0.74032 | -82.6888
0.74063 | -82.7134
0.74094 | -82.7385
0.74125 | -82.7632
0.74156 | -82.7862
0.74187 | -82.8157
0.74218 | -82.8352
0.7425 | -82.8581
0.74281 | -82.8957
0.74312 | -82.9409
0.74343 | -82.965
0.74374 | -82.9886
0.74405 | -83.0132
0.74436 | -83.0333
0.74467 | -83.0582
0.74498 | -83.0826
0.74529 | -83.1091
0.7456 | -83.1372
0.74591 | -83.1569
0.74622 | -83.1817
0.74653 | -83.203
0.74684 | -83.2203
0.74715 | -83.2408
0.74746 | -83.2619
0.74777 | -83.2862
0.74809 | -83.3096
0.7484 | -83.334
0.74871 | -83.3598
0.74902 | -83.385
0.74933 | -83.4059
0.74964 | -83.4263
0.74995 | -83.4464
0.75026 | -83.4685
0.75057 | -83.4926
0.75088 | -83.5147
0.75119 | -83.5367
0.7515 | -83.5618
0.75181 | -83.5801
0.75212 | -83.5923
0.75243 | -83.6031
0.75274 | -83.6244
0.75305 | -83.6436
0.75336 | -83.6707
0.75367 | -83.6845
0.75399 | -83.6998
0.7543 | -83.7126
0.75461 | -83.7233
0.75492 | -83.7385
0.75523 | -83.748
0.75554 | -83.7662
0.75585 | -83.7734
0.75616 | -83.7851
0.75647 | -83.7972
0.75678 | -83.8085
0.75709 | -83.8273
0.7574 | -83.8365
0.75771 | -83.8402
0.75802 | -83.8468
0.75833 | -83.8536
0.75864 | -83.8659
0.75895 | -83.8728
0.75926 | -83.8806
0.75958 | -83.8809
0.75989 | -83.8854
0.7602 | -83.8896
0.76051 | -83.894
0.76082 | -83.8967
0.76113 | -83.9049
0.76144 | -83.9087
0.76175 | -83.9145
0.76206 | -83.9189
0.76237 | -83.9233
0.76268 | -83.9219
0.76299 | -83.9295
0.7633 | -83.9213
0.76361 | -83.9224
0.76392 | -83.9229
0.76423 | -83.9333
0.76454 | -83.9312
0.76485 | -83.9281
0.76516 | -83.9377
0.76548 | -83.9387
0.76579 | -83.9397
0.7661 | -83.9348
0.76641 | -83.9364
0.76672 | -83.9367
0.76703 | -83.9365
0.76734 | -83.9386
0.76765 | -83.938
0.76796 | -83.9433
0.76827 | -83.937
0.76858 | -83.9342
0.76889 | -83.9308
0.7692 | -83.9266
0.76951 | -83.9291
0.76982 | -83.9317
0.77013 | -83.9217
0.77044 | -83.9062
0.77075 | -83.9028
0.77107 | -83.9059
0.77138 | -83.9016
0.77169 | -83.8985
0.772 | -83.8944
0.77231 | -83.8879
0.77262 | -83.8823
0.77293 | -83.8789
0.77324 | -83.8661
0.77355 | -83.8614
0.77386 | -83.8647
0.77417 | -83.8528
0.77448 | -83.8391
0.77479 | -83.8349
0.7751 | -83.8286
0.77541 | -83.8328
0.77572 | -83.8253
0.77603 | -83.8211
0.77634 | -83.8085
0.77665 | -83.8013
0.77697 | -83.7969
0.77728 | -83.7908
0.77759 | -83.7858
0.7779 | -83.7805
0.77821 | -83.7679
0.77852 | -83.7519
0.77883 | -83.7439
0.77914 | -83.734
0.77945 | -83.7244
0.77976 | -83.7149
0.78007 | -83.7006
0.78038 | -83.6924
0.78069 | -83.6838
0.781 | -83.6763
0.78131 | -83.6703
0.78162 | -83.6622
0.78193 | -83.6524
0.78224 | -83.6383
0.78256 | -83.6284
0.78287 | -83.6191
0.78318 | -83.6092
0.78349 | -83.5964
0.7838 | -83.5828
0.78411 | -83.5737
0.78442 | -83.5623
0.78473 | -83.5473
0.78504 | -83.538
0.78535 | -83.5296
0.78566 | -83.5174
0.78597 | -83.503
0.78628 | -83.4926
0.78659 | -83.4815
0.7869 | -83.4681
0.78721 | -83.4615
0.78752 | -83.4497
0.78783 | -83.4367
0.78814 | -83.4258
0.78846 | -83.4075
0.78877 | -83.3943
0.78908 | -83.3796
0.78939 | -83.3628
0.7897 | -83.3294
0.79001 | -83.3117
0.79032 | -83.2954
0.79063 | -83.2803
0.79094 | -83.2754
0.79125 | -83.2578
0.79156 | -83.2427
0.79187 | -83.2285
0.79218 | -83.2181
0.79249 | -83.2055
0.7928 | -83.1906
0.79311 | -83.179
0.79342 | -83.1647
0.79373 | -83.1511
0.79405 | -83.1366
0.79436 | -83.1219
0.79467 | -83.1086
0.79498 | -83.0934
0.79529 | -83.0841
0.7956 | -83.0727
0.79591 | -83.0636
0.79622 | -83.0519
0.79653 | -83.0387
0.79684 | -83.0255
0.79715 | -83.0196
0.79746 | -83.0076
0.79777 | -83
0.79808 | -82.9868
0.79839 | -82.9768
0.7987 | -82.9656
0.79901 | -82.957
0.79932 | -82.9475
0.79963 | -82.9304
0.79995 | -82.9209
0.80026 | -82.9127
0.80057 | -82.9025
0.80088 | -82.8957
0.80119 | -82.8919
0.8015 | -82.8851
0.80181 | -82.8785
0.80212 | -82.8746
0.80243 | -82.8709
0.80274 | -82.8648
0.80305 | -82.856
0.80336 | -82.8495
0.80367 | -82.8439
0.80398 | -82.837
0.80429 | -82.8319
0.8046 | -82.8288
0.80491 | -82.8221
0.80522 | -82.817
0.80554 | -82.8137
0.80585 | -82.8079
0.80616 | -82.8124
0.80647 | -82.8093
0.80678 | -82.8111
0.80709 | -82.8119
0.8074 | -82.8117
0.80771 | -82.813
0.80802 | -82.8146
0.80833 | -82.8167
0.80864 | -82.8187
0.80895 | -82.8205
0.80926 | -82.8237
0.80957 | -82.8282
0.80988 | -82.8306
0.81019 | -82.8349
0.8105 | -82.8406
0.81081 | -82.8468
0.81112 | -82.8556
0.81144 | -82.8614
0.81175 | -82.8665
0.81206 | -82.8744
0.81237 | -82.8855
0.81268 | -82.8955
0.81299 | -82.9046
0.8133 | -82.9156
0.81361 | -82.9283
0.81392 | -82.9388
0.81423 | -82.947
0.81454 | -82.9614
0.81485 | -82.9715
0.81516 | -82.9862
0.81547 | -82.9986
0.81578 | -83.0098
0.81609 | -83.0198
0.8164 | -83.0286
0.81671 | -83.0367
0.81703 | -83.0488
0.81734 | -83.0652
0.81765 | -83.081
0.81796 | -83.0972
0.81827 | -83.1186
0.81858 | -83.1324
0.81889 | -83.1578
0.8192 | -83.1824
0.81951 | -83.1993
0.81982 | -83.2181
0.82013 | -83.2392
0.82044 | -83.2603
0.82075 | -83.2786
0.82106 | -83.2992
0.82137 | -83.3177
0.82168 | -83.3357
0.82199 | -83.3532
0.8223 | -83.3735
0.82261 | -83.3957
0.82293 | -83.418
0.82324 | -83.4442
0.82355 | -83.4728
0.82386 | -83.4976
0.82417 | -83.5228
0.82448 | -83.5552
0.82479 | -83.581
0.8251 | -83.6069
0.82541 | -83.6309
0.82572 | -83.6495
0.82603 | -83.6699
0.82634 | -83.6958
0.82665 | -83.719
0.82696 | -83.7412
0.82727 | -83.7564
0.82758 | -83.7776
0.82789 | -83.799
0.8282 | -83.8229
0.82852 | -83.8469
0.82883 | -83.8696
0.82914 | -83.8868
0.82945 | -83.9104
0.82976 | -83.9359
0.83007 | -83.956
0.83038 | -83.9809
0.83069 | -84.0004
0.831 | -84.0139
0.83131 | -84.0311
0.83162 | -84.0476
0.83193 | -84.0671
0.83224 | -84.077
0.83255 | -84.0923
0.83286 | -84.1006
0.83317 | -84.104
0.83348 | -84.1149
0.83379 | -84.1268
0.8341 | -84.1393
0.83442 | -84.1513
0.83473 | -84.1691
0.83504 | -84.1889
0.83535 | -84.1992
0.83566 | -84.2125
0.83597 | -84.2252
0.83628 | -84.2328
0.83659 | -84.2409
0.8369 | -84.2562
0.83721 | -84.2631
0.83752 | -84.2705
0.83783 | -84.2776
0.83814 | -84.2778
0.83845 | -84.2838
0.83876 | -84.2929
0.83907 | -84.2938
0.83938 | -84.2839
0.83969 | -84.2734
0.84001 | -84.2707
0.84032 | -84.272
0.84063 | -84.2733
0.84094 | -84.2701
0.84125 | -84.2694
0.84156 | -84.2674
0.84187 | -84.2616
0.84218 | -84.2473
0.84249 | -84.2415
0.8428 | -84.2363
0.84311 | -84.2342
0.84342 | -84.2278
0.84373 | -84.2214
0.84404 | -84.2165
0.84435 | -84.2092
0.84466 | -84.2101
0.84497 | -84.2067
0.84528 | -84.1968
0.84559 | -84.1933
0.84591 | -84.1888
0.84622 | -84.1766
0.84653 | -84.1706
0.84684 | -84.1655
0.84715 | -84.1824
0.84746 | -84.1853
0.84777 | -84.1754
0.84808 | -84.1659
0.84839 | -84.1523
0.8487 | -84.1416
0.84901 | -84.1293
0.84932 | -84.1135
0.84963 | -84.1014
0.84994 | -84.0906
0.85025 | -84.0828
0.85056 | -84.0732
0.85087 | -84.07
0.85118 | -84.0602
0.8515 | -84.0442
0.85181 | -84.0347
0.85212 | -84.0246
0.85243 | -84.0139
0.85274 | -84.0068
0.85305 | -83.9896
0.85336 | -83.9745
0.85367 | -83.9677
0.85398 | -83.962
0.85429 | -83.9543
0.8546 | -83.9412
0.85491 | -83.9319
0.85522 | -83.9261
0.85553 | -83.9184
0.85584 | -83.9141
0.85615 | -83.9149
0.85646 | -83.9062
0.85677 | -83.9037
0.85708 | -83.9
0.8574 | -83.895
0.85771 | -83.9
0.85802 | -83.8907
0.85833 | -83.8848
0.85864 | -83.8834
0.85895 | -83.8797
0.85926 | -83.8781
0.85957 | -83.8705
0.85988 | -83.8679
0.86019 | -83.8626
0.8605 | -83.8623
0.86081 | -83.8608
0.86112 | -83.8555
0.86143 | -83.8517
0.86174 | -83.8513
0.86205 | -83.8497
0.86236 | -83.8508
0.86267 | -83.8512
0.86299 | -83.8518
0.8633 | -83.8504
0.86361 | -83.8526
0.86392 | -83.8521
0.86423 | -83.857
0.86454 | -83.861
0.86485 | -83.8634
0.86516 | -83.8654
0.86547 | -83.8679
0.86578 | -83.8684
0.86609 | -83.8699
0.8664 | -83.874
0.86671 | -83.8754
0.86702 | -83.879
0.86733 | -83.8848
0.86764 | -83.8853
0.86795 | -83.8897
0.86826 | -83.8925
0.86857 | -83.8946
0.86889 | -83.8996
0.8692 | -83.9009
0.86951 | -83.8919
0.86982 | -83.8923
0.87013 | -83.8958
0.87044 | -83.8993
0.87075 | -83.9036
0.87106 | -83.907
0.87137 | -83.9053
0.87168 | -83.9087
0.87199 | -83.9064
0.8723 | -83.9015
0.87261 | -83.9036
0.87292 | -83.9065
0.87323 | -83.9123
0.87354 | -83.9098
0.87385 | -83.9137
0.87416 | -83.9187
0.87447 | -83.9205
0.87479 | -83.9206
0.87509 | -83.9238
0.87541 | -83.9265
0.87572 | -83.9303
0.87603 | -83.9338
0.87634 | -83.932
0.87665 | -83.931
0.87696 | -83.9295
0.87727 | -83.9289
0.87758 | -83.9284
0.87789 | -83.9269
0.8782 | -83.9283
0.87851 | -83.9278
0.87882 | -83.9277
0.87913 | -83.9277
0.87944 | -83.9255
0.87975 | -83.9196
0.88006 | -83.9148
0.88038 | -83.9083
0.88069 | -83.9023
0.881 | -83.9029
0.88131 | -83.9059
0.88162 | -83.9067
0.88193 | -83.9046
0.88224 | -83.9004
0.88255 | -83.8968
0.88286 | -83.8921
0.88317 | -83.8889
0.88348 | -83.8875
0.88379 | -83.887
0.8841 | -83.8842
0.88441 | -83.8803
0.88472 | -83.8752
0.88503 | -83.8726
0.88534 | -83.8656
0.88565 | -83.8649
0.88596 | -83.8721
0.88628 | -83.8755
0.88658 | -83.8732
0.8869 | -83.8731
0.88721 | -83.869
0.88752 | -83.8722
0.88783 | -83.8687
0.88814 | -83.86
0.88845 | -83.8494
0.88876 | -83.8465
0.88907 | -83.8448
0.88938 | -83.8462
0.88969 | -83.8483
0.89 | -83.8419
0.89031 | -83.8373
0.89062 | -83.8363
0.89093 | -83.8422
0.89124 | -83.8461
0.89155 | -83.8469
0.89187 | -83.8462
0.89218 | -83.8485
0.89249 | -83.8525
0.8928 | -83.8525
0.89311 | -83.8611
0.89342 | -83.8604
0.89373 | -83.863
0.89404 | -83.8661
0.89435 | -83.8654
0.89466 | -83.864
0.89497 | -83.8649
0.89528 | -83.8674
0.89559 | -83.8685
0.8959 | -83.8738
0.89621 | -83.8737
0.89652 | -83.8761
0.89683 | -83.8782
0.89714 | -83.8808
0.89745 | -83.8868
0.89777 | -83.8905
0.89807 | -83.8961
0.89839 | -83.9015
0.8987 | -83.9038
0.89901 | -83.9143
0.89932 | -83.9229
0.89963 | -83.9329
0.89994 | -83.9474
0.90025 | -83.9658
0.90056 | -83.9846
0.90087 | -83.9951
0.90118 | -84.0043
0.90149 | -84.014
0.9018 | -84.0217
0.90211 | -84.0318
0.90242 | -84.0451
0.90273 | -84.0566
0.90304 | -84.068
0.90335 | -84.0803
0.90367 | -84.0883
0.90398 | -84.101
0.90429 | -84.1135
0.9046 | -84.1265
0.90491 | -84.1462
0.90522 | -84.1592
0.90553 | -84.1738
0.90584 | -84.1852
0.90615 | -84.1983
0.90646 | -84.2107
0.90677 | -84.2277
0.90708 | -84.2434
0.90739 | -84.2619
0.9077 | -84.2825
0.90801 | -84.2921
0.90832 | -84.3053
0.90863 | -84.3218
0.90894 | -84.3298
0.90926 | -84.3395
0.90956 | -84.3547
0.90988 | -84.3725
0.91019 | -84.3916
0.9105 | -84.409
0.91081 | -84.4251
0.91112 | -84.4363
0.91143 | -84.4518
0.91174 | -84.4677
0.91205 | -84.4817
0.91236 | -84.4968
0.91267 | -84.512
0.91298 | -84.5306
0.91329 | -84.554
0.9136 | -84.5714
0.91391 | -84.5847
0.91422 | -84.6009
0.91453 | -84.621
0.91484 | -84.6394
0.91516 | -84.6651
0.91546 | -84.6862
0.91578 | -84.6988
0.91609 | -84.7149
0.9164 | -84.7315
0.91671 | -84.753
0.91702 | -84.7698
0.91733 | -84.7855
0.91764 | -84.7998
0.91795 | -84.8137
0.91826 | -84.8208
0.91857 | -84.8377
0.91888 | -84.8521
0.91919 | -84.8668
0.9195 | -84.8796
0.91981 | -84.8924
0.92012 | -84.9108
0.92043 | -84.9222
0.92075 | -84.9325
0.92105 | -84.9438
0.92137 | -84.9543
0.92168 | -84.9726
0.92199 | -84.9817
0.9223 | -85.0038
0.92261 | -85.0093
0.92292 | -85.0217
0.92323 | -85.0267
0.92354 | -85.0264
0.92385 | -85.035
0.92416 | -85.0505
0.92447 | -85.0582
0.92478 | -85.0665
0.92509 | -85.0753
0.9254 | -85.0836
0.92571 | -85.0876
0.92602 | -85.1016
0.92633 | -85.1063
0.92665 | -85.111
0.92695 | -85.1178
0.92727 | -85.1229
0.92758 | -85.1277
0.92789 | -85.1365
0.9282 | -85.1403
0.92851 | -85.1427
0.92882 | -85.1449
0.92913 | -85.147
0.92944 | -85.1473
0.92975 | -85.1493
0.93006 | -85.1497
0.93037 | -85.1555
0.93068 | -85.1576
0.93099 | -85.1613
0.9313 | -85.1645
0.93161 | -85.1643
0.93192 | -85.1639
0.93223 | -85.157
0.93254 | -85.1618
0.93286 | -85.1799
0.93317 | -85.1827
0.93348 | -85.1808
0.93379 | -85.1781
0.9341 | -85.1758
0.93441 | -85.1734
0.93472 | -85.1769
0.93503 | -85.1747
0.93534 | -85.1717
0.93565 | -85.1718
0.93596 | -85.1659
0.93627 | -85.1622
0.93658 | -85.1536
0.93689 | -85.1466
0.9372 | -85.1445
0.93751 | -85.1407
0.93782 | -85.1395
0.93814 | -85.1348
0.93844 | -85.1269
0.93876 | -85.1247
0.93907 | -85.1197
0.93938 | -85.1121
0.93969 | -85.1088
0.94 | -85.1014
0.94031 | -85.0885
0.94062 | -85.0821
0.94093 | -85.0745
0.94124 | -85.0695
0.94155 | -85.0641
0.94186 | -85.0576
0.94217 | -85.0514
0.94248 | -85.0458
0.94279 | -85.0375
0.9431 | -85.028
0.94341 | -85.0232
0.94372 | -85.0225
0.94403 | -85.0204
0.94434 | -85.022
0.94466 | -85.0201
0.94497 | -85.0138
0.94528 | -85.0091
0.94559 | -85.0035
0.9459 | -84.9942
0.94621 | -84.9878
0.94652 | -84.9831
0.94683 | -84.9799
0.94714 | -84.9753
0.94745 | -84.9697
0.94776 | -84.9612
0.94807 | -84.9572
0.94838 | -84.948
0.94869 | -84.9423
0.949 | -84.9354
0.94931 | -84.9373
0.94963 | -84.9333
0.94993 | -84.9332
0.95025 | -84.931
0.95056 | -84.9269
0.95087 | -84.9221
0.95118 | -84.9169
0.95149 | -84.9137
0.9518 | -84.9098
0.95211 | -84.9079
0.95242 | -84.9079
0.95273 | -84.9121
0.95304 | -84.9149
0.95335 | -84.9132
0.95366 | -84.9115
0.95397 | -84.9105
0.95428 | -84.9112
0.95459 | -84.9141
0.9549 | -84.9147
0.95521 | -84.9162
0.95552 | -84.9165
0.95583 | -84.9158
0.95615 | -84.9162
0.95646 | -84.9191
0.95677 | -84.9194
0.95708 | -84.9207
0.95739 | -84.9295
0.9577 | -84.9315
0.95801 | -84.9361
0.95832 | -84.9384
0.95863 | -84.9486
0.95894 | -84.9527
0.95925 | -84.96
0.95956 | -84.9616
0.95987 | -84.9601
0.96018 | -84.959
0.96049 | -84.9588
0.9608 | -84.9643
0.96111 | -84.9736
0.96142 | -84.9831
0.96174 | -84.9863
0.96205 | -84.9905
0.96236 | -84.9977
0.96267 | -85.002
0.96298 | -85.0095
0.96329 | -85.0156
0.9636 | -85.0209
0.96391 | -85.0318
0.96422 | -85.0383
0.96453 | -85.0454
0.96484 | -85.0518
0.96515 | -85.0558
0.96546 | -85.059
0.96577 | -85.0658
0.96608 | -85.0687
0.96639 | -85.0521
0.9667 | -85.051
0.96701 | -85.0594
0.96732 | -85.0658
0.96764 | -85.0566
0.96795 | -85.0545
0.96826 | -85.0594
0.96857 | -85.0649
0.96888 | -85.075
0.96919 | -85.0866
0.9695 | -85.0894
0.96981 | -85.0936
0.97012 | -85.0935
0.97043 | -85.0903
0.97074 | -85.0969
0.97105 | -85.1032
0.97136 | -85.1054
0.97167 | -85.0985
0.97198 | -85.0943
0.97229 | -85.0994
0.9726 | -85.1
0.97291 | -85.098
0.97323 | -85.1
0.97354 | -85.1044
0.97385 | -85.1124
0.97416 | -85.1179
0.97447 | -85.1188
0.97478 | -85.128
0.97509 | -85.1365
0.9754 | -85.1381
0.97571 | -85.1372
0.97602 | -85.1329
0.97633 | -85.1203
0.97664 | -85.1128
0.97695 | -85.1145
0.97726 | -85.1175
0.97757 | -85.1161
0.97788 | -85.113
0.97819 | -85.1057
0.9785 | -85.1024
0.97881 | -85.0969
0.97913 | -85.0886
0.97944 | -85.0827
0.97975 | -85.0768
0.98006 | -85.0717
0.98037 | -85.0697
0.98068 | -85.0656
0.98099 | -85.0593
0.9813 | -85.0525
0.98161 | -85.0423
0.98192 | -85.0316
0.98223 | -85.0218
0.98254 | -85.0101
0.98285 | -84.9993
0.98316 | -84.9838
0.98347 | -84.9738
0.98378 | -84.9642
0.98409 | -84.9576
0.9844 | -84.9463
0.98472 | -84.9435
0.98503 | -84.9336
0.98534 | -84.921
0.98565 | -84.911
0.98596 | -84.8999
0.98627 | -84.8874
0.98658 | -84.8776
0.98689 | -84.8677
0.9872 | -84.8589
0.98751 | -84.8486
0.98782 | -84.8384
0.98813 | -84.8276
0.98844 | -84.8125
0.98875 | -84.8031
0.98906 | -84.7899
0.98937 | -84.778
0.98968 | -84.7684
0.98999 | -84.7578
0.9903 | -84.7498
0.99062 | -84.7462
0.99093 | -84.7381
0.99124 | -84.726
0.99155 | -84.7141
0.99186 | -84.7018
0.99217 | -84.6889
0.99248 | -84.6716
0.99279 | -84.6595
0.9931 | -84.6488
0.99341 | -84.6407
0.99372 | -84.6321
0.99403 | -84.6198
0.99434 | -84.6156
0.99465 | -84.6029
0.99496 | -84.5907
0.99527 | -84.5786
0.99558 | -84.5684
0.99589 | -84.5635
0.99621 | -84.5565
0.99652 | -84.5491
0.99683 | -84.5458
0.99714 | -84.5369
0.99745 | -84.5277
0.99776 | -84.5161
0.99807 | -84.5086
0.99838 | -84.5019
0.99869 | -84.4976
0.999 | -84.4901
0.99931 | -84.4849
0.99962 | -84.4838
0.99993 | -84.4864
1.0002 | -84.4839
1.0006 | -84.4793
1.0009 | -84.4755
1.0012 | -84.4688
1.0015 | -84.4613
1.0018 | -84.4596
1.0021 | -84.4584
1.0024 | -84.4582
1.0027 | -84.4578
1.003 | -84.4635
1.0033 | -84.4741
1.0037 | -84.4805
1.004 | -84.4794
1.0043 | -84.4761
1.0046 | -84.4813
1.0049 | -84.4906
1.0052 | -84.493
1.0055 | -84.4956
1.0058 | -84.4967
1.0061 | -84.4976
1.0065 | -84.4999
1.0068 | -84.5047
1.0071 | -84.5093
1.0074 | -84.5156
1.0077 | -84.5221
1.008 | -84.5286
1.0083 | -84.537
1.0086 | -84.5405
1.0089 | -84.5462
1.0092 | -84.55
1.0096 | -84.5538
1.0099 | -84.5612
1.0102 | -84.5672
1.0105 | -84.5686
1.0108 | -84.5674
1.0111 | -84.5719
1.0114 | -84.5822
1.0117 | -84.5878
1.012 | -84.5937
1.0124 | -84.5885
1.0127 | -84.5848
1.013 | -84.591
1.0133 | -84.5974
1.0136 | -84.6033
1.0139 | -84.6111
1.0142 | -84.6136
1.0145 | -84.6207
1.0148 | -84.6218
1.0151 | -84.6291
1.0155 | -84.6368
1.0158 | -84.6441
1.0161 | -84.6533
1.0164 | -84.6632
1.0167 | -84.6686
1.017 | -84.6815
1.0173 | -84.6912
1.0176 | -84.6983
1.0179 | -84.7049
1.0183 | -84.7119
1.0186 | -84.7157
1.0189 | -84.7195
1.0192 | -84.7215
1.0195 | -84.7237
1.0198 | -84.7301
1.0201 | -84.7367
1.0204 | -84.7437
1.0207 | -84.7492
1.021 | -84.7477
1.0214 | -84.7473
1.0217 | -84.7469
1.022 | -84.7488
1.0223 | -84.7551
1.0226 | -84.7552
1.0229 | -84.7551
1.0232 | -84.7527
1.0235 | -84.7536
1.0238 | -84.7545
1.0242 | -84.7573
1.0245 | -84.755
1.0248 | -84.7548
1.0251 | -84.7531
1.0254 | -84.7525
1.0257 | -84.7509
1.026 | -84.7507
1.0263 | -84.751
1.0266 | -84.748
1.0269 | -84.7428
1.0273 | -84.7381
1.0276 | -84.7364
1.0279 | -84.7318
1.0282 | -84.728
1.0285 | -84.7251
1.0288 | -84.7188
1.0291 | -84.7105
1.0294 | -84.7102
1.0297 | -84.7092
1.0301 | -84.7079
1.0304 | -84.6991
1.0307 | -84.6939
1.031 | -84.6868
1.0313 | -84.6783
1.0316 | -84.674
1.0319 | -84.6676
1.0322 | -84.6605
1.0325 | -84.6535
1.0328 | -84.6463
1.0332 | -84.6387
1.0335 | -84.6325
1.0338 | -84.6236
1.0341 | -84.6097
1.0344 | -84.6019
1.0347 | -84.5875
1.035 | -84.5758
1.0353 | -84.5672
1.0356 | -84.558
1.036 | -84.5528
1.0363 | -84.5635
1.0366 | -84.5575
1.0369 | -84.5444
1.0372 | -84.5347
1.0375 | -84.5274
1.0378 | -84.529
1.0381 | -84.5256
1.0384 | -84.5169
1.0387 | -84.5077
1.0391 | -84.5038
1.0394 | -84.4936
1.0397 | -84.4843
1.04 | -84.4757
1.0403 | -84.4662
1.0406 | -84.4611
1.0409 | -84.4597
1.0412 | -84.4504
1.0415 | -84.4408
1.0419 | -84.4318
1.0422 | -84.4283
1.0425 | -84.4251
1.0428 | -84.4162
1.0431 | -84.4105
1.0434 | -84.4068
1.0437 | -84.3987
1.044 | -84.3907
1.0443 | -84.3854
1.0446 | -84.3784
1.045 | -84.3715
1.0453 | -84.3669
1.0456 | -84.3629
1.0459 | -84.3582
1.0462 | -84.3511
1.0465 | -84.3488
1.0468 | -84.3522
1.0471 | -84.3489
1.0474 | -84.3508
1.0478 | -84.3479
1.0481 | -84.3563
1.0484 | -84.3555
1.0487 | -84.3558
1.049 | -84.3556
1.0493 | -84.3601
1.0496 | -84.3593
1.0499 | -84.3584
1.0502 | -84.3579
1.0505 | -84.3589
1.0509 | -84.361
1.0512 | -84.3635
1.0515 | -84.3663
1.0518 | -84.3703
1.0521 | -84.3727
1.0524 | -84.3741
1.0527 | -84.3809
1.053 | -84.3857
1.0533 | -84.3933
1.0537 | -84.4009
1.054 | -84.4086
1.0543 | -84.42
1.0546 | -84.4244
1.0549 | -84.431
1.0552 | -84.4372
1.0555 | -84.442
1.0558 | -84.4478
1.0561 | -84.4558
1.0564 | -84.4634
1.0568 | -84.4701
1.0571 | -84.4834
1.0574 | -84.4935
1.0577 | -84.5026
1.058 | -84.5162
1.0583 | -84.5251
1.0586 | -84.5343
1.0589 | -84.5448
1.0592 | -84.5552
1.0596 | -84.5631
1.0599 | -84.5725
1.0602 | -84.5859
1.0605 | -84.5969
1.0608 | -84.6115
1.0611 | -84.6246
1.0614 | -84.6378
1.0617 | -84.6512
1.062 | -84.662
1.0623 | -84.6719
1.0627 | -84.6805
1.063 | -84.6869
1.0633 | -84.6975
1.0636 | -84.7122
1.0639 | -84.7239
1.0642 | -84.7376
1.0645 | -84.7521
1.0648 | -84.7666
1.0651 | -84.7798
1.0655 | -84.7953
1.0658 | -84.8121
1.0661 | -84.8256
1.0664 | -84.835
1.0667 | -84.8505
1.067 | -84.8668
1.0673 | -84.8814
1.0676 | -84.8958
1.0679 | -84.9104
1.0682 | -84.9248
1.0686 | -84.9394
1.0689 | -84.9545
1.0692 | -84.9676
1.0695 | -84.9764
1.0698 | -84.987
1.0701 | -85.0024
1.0704 | -85.0159
1.0707 | -85.0296
1.071 | -85.0433
1.0714 | -85.0531
1.0717 | -85.0606
1.072 | -85.072
1.0723 | -85.0853
1.0726 | -85.0982
1.0729 | -85.1125
1.0732 | -85.1265
1.0735 | -85.1409
1.0738 | -85.1562
1.0741 | -85.1644
1.0745 | -85.1754
1.0748 | -85.187
1.0751 | -85.1967
1.0754 | -85.2082
1.0757 | -85.2199
1.076 | -85.231
1.0763 | -85.2383
1.0766 | -85.2467
1.0769 | -85.2617
1.0773 | -85.2745
1.0776 | -85.2845
1.0779 | -85.2949
1.0782 | -85.3048
1.0785 | -85.316
1.0788 | -85.3285
1.0791 | -85.3386
1.0794 | -85.3504
1.0797 | -85.3586
1.08 | -85.3687
1.0804 | -85.3793
1.0807 | -85.388
1.081 | -85.3961
1.0813 | -85.4044
1.0816 | -85.4145
1.0819 | -85.4216
1.0822 | -85.427
1.0825 | -85.4355
1.0828 | -85.4417
1.0832 | -85.4454
1.0835 | -85.4532
1.0838 | -85.4672
1.0841 | -85.4688
1.0844 | -85.4778
1.0847 | -85.4825
1.085 | -85.4884
1.0853 | -85.4937
1.0856 | -85.509
1.086 | -85.5191
1.0863 | -85.5309
1.0866 | -85.5263
1.0869 | -85.5229
1.0872 | -85.5255
1.0875 | -85.5269
1.0878 | -85.5285
1.0881 | -85.5242
1.0884 | -85.5308
1.0887 | -85.5299
1.0891 | -85.526
1.0894 | -85.5182
1.0897 | -85.5148
1.09 | -85.5139
1.0903 | -85.5142
1.0906 | -85.5123
1.0909 | -85.5123
1.0912 | -85.5122
1.0915 | -85.5086
1.0919 | -85.5054
1.0922 | -85.5026
1.0925 | -85.4931
1.0928 | -85.4859
1.0931 | -85.4862
1.0934 | -85.483
1.0937 | -85.4767
1.094 | -85.4717
1.0943 | -85.4681
1.0946 | -85.4653
1.095 | -85.467
1.0953 | -85.4622
1.0956 | -85.4571
1.0959 | -85.4558
1.0962 | -85.4503
1.0965 | -85.446
1.0968 | -85.4429
1.0971 | -85.4389
1.0974 | -85.4377
1.0978 | -85.4321
1.0981 | -85.4276
1.0984 | -85.423
1.0987 | -85.4211
1.099 | -85.4136
1.0993 | -85.4088
1.0996 | -85.4042
1.0999 | -85.4003
1.1002 | -85.3948
1.1005 | -85.386
1.1009 | -85.3818
1.1012 | -85.3806
1.1015 | -85.3775
1.1018 | -85.3743
1.1021 | -85.3699
1.1024 | -85.3657
1.1027 | -85.3615
1.103 | -85.3547
1.1033 | -85.3508
1.1037 | -85.3544
1.104 | -85.3515
1.1043 | -85.3479
1.1046 | -85.3435
1.1049 | -85.3389
1.1052 | -85.3353
1.1055 | -85.3301
1.1058 | -85.3296
1.1061 | -85.3279
1.1064 | -85.3221
1.1068 | -85.3163
1.1071 | -85.3133
1.1074 | -85.3114
1.1077 | -85.309
1.108 | -85.3028
1.1083 | -85.2912
1.1086 | -85.2864
1.1089 | -85.2847
1.1092 | -85.2796
1.1096 | -85.2843
1.1099 | -85.2862
1.1102 | -85.2869
1.1105 | -85.2842
1.1108 | -85.2868
1.1111 | -85.2852
1.1114 | -85.2813
1.1117 | -85.2777
1.112 | -85.2751
1.1123 | -85.2706
1.1127 | -85.2685
1.113 | -85.2648
1.1133 | -85.2585
1.1136 | -85.2488
1.1139 | -85.2436
1.1142 | -85.2398
1.1145 | -85.2321
1.1148 | -85.2196
1.1151 | -85.2118
1.1155 | -85.2075
1.1158 | -85.2052
1.1161 | -85.2035
1.1164 | -85.1999
1.1167 | -85.1984
1.117 | -85.1951
1.1173 | -85.1958
1.1176 | -85.2006
1.1179 | -85.1992
1.1182 | -85.1928
1.1186 | -85.189
1.1189 | -85.1816
1.1192 | -85.1777
1.1195 | -85.1763
1.1198 | -85.1774
1.1201 | -85.1748
1.1204 | -85.1727
1.1207 | -85.1717
1.121 | -85.1688
1.1214 | -85.1656
1.1217 | -85.163
1.122 | -85.1649
1.1223 | -85.1626
1.1226 | -85.1597
1.1229 | -85.1582
1.1232 | -85.1564
1.1235 | -85.156
1.1238 | -85.154
1.1241 | -85.1507
1.1245 | -85.1465
1.1248 | -85.1449
1.1251 | -85.1433
1.1254 | -85.1417
1.1257 | -85.1423
1.126 | -85.1444
1.1263 | -85.1439
1.1266 | -85.1407
1.1269 | -85.139
1.1273 | -85.1364
1.1276 | -85.135
1.1279 | -85.1333
1.1282 | -85.1332
1.1285 | -85.134
1.1288 | -85.1269
1.1291 | -85.1247
1.1294 | -85.1267
1.1297 | -85.1247
1.13 | -85.1196
1.1304 | -85.1139
1.1307 | -85.1137
1.131 | -85.1108
1.1313 | -85.1079
1.1316 | -85.1066
1.1319 | -85.1094
1.1322 | -85.1083
1.1325 | -85.1064
1.1328 | -85.1063
1.1332 | -85.1071
1.1335 | -85.1088
1.1338 | -85.1078
1.1341 | -85.1048
1.1344 | -85.1046
1.1347 | -85.104
1.135 | -85.1039
1.1353 | -85.1004
1.1356 | -85.1
1.1359 | -85.1007
1.1363 | -85.1021
1.1366 | -85.1043
1.1369 | -85.1033
1.1372 | -85.1034
1.1375 | -85.1031
1.1378 | -85.1067
1.1381 | -85.1066
1.1384 | -85.1069
1.1387 | -85.107
1.1391 | -85.1092
1.1394 | -85.1125
1.1397 | -85.1163
1.14 | -85.1221
1.1403 | -85.1259
1.1406 | -85.1255
1.1409 | -85.1279
1.1412 | -85.1339
1.1415 | -85.136
1.1418 | -85.1363
1.1422 | -85.1372
1.1425 | -85.1386
1.1428 | -85.1387
1.1431 | -85.1416
1.1434 | -85.1434
1.1437 | -85.1424
1.144 | -85.1445
1.1443 | -85.1462
1.1446 | -85.1505
1.145 | -85.1553
1.1453 | -85.1598
1.1456 | -85.1623
1.1459 | -85.1656
1.1462 | -85.1674
1.1465 | -85.1717
1.1468 | -85.1822
1.1471 | -85.1899
1.1474 | -85.2007
1.1477 | -85.2071
1.1481 | -85.211
1.1484 | -85.2144
1.1487 | -85.2188
1.149 | -85.2291
1.1493 | -85.2369
1.1496 | -85.2418
1.1499 | -85.2452
1.1502 | -85.2499
1.1505 | -85.2554
1.1509 | -85.2596
1.1512 | -85.2677
1.1515 | -85.2765
1.1518 | -85.2863
1.1521 | -85.293
1.1524 | -85.3009
1.1527 | -85.3093
1.153 | -85.3158
1.1533 | -85.3235
1.1536 | -85.3316
1.154 | -85.341
1.1543 | -85.3524
1.1546 | -85.3632
1.1549 | -85.3698
1.1552 | -85.3763
1.1555 | -85.3901
1.1558 | -85.3983
1.1561 | -85.408
1.1564 | -85.4162
1.1568 | -85.4265
1.1571 | -85.4373
1.1574 | -85.4498
1.1577 | -85.4617
1.158 | -85.4745
1.1583 | -85.4865
1.1586 | -85.4963
1.1589 | -85.5027
1.1592 | -85.5141
1.1595 | -85.5286
1.1599 | -85.5415
1.1602 | -85.5561
1.1605 | -85.574
1.1608 | -85.5812
1.1611 | -85.588
1.1614 | -85.6029
1.1617 | -85.6213
1.162 | -85.6328
1.1623 | -85.6454
1.1627 | -85.6583
1.163 | -85.672
1.1633 | -85.684
1.1636 | -85.6973
1.1639 | -85.7116
1.1642 | -85.7257
1.1645 | -85.7414
1.1648 | -85.7547
1.1651 | -85.7713
1.1654 | -85.7907
1.1658 | -85.8081
1.1661 | -85.8166
1.1664 | -85.8296
1.1667 | -85.8446
1.167 | -85.8629
1.1673 | -85.8776
1.1676 | -85.8897
1.1679 | -85.904
1.1682 | -85.9232
1.1686 | -85.9368
1.1689 | -85.9495
1.1692 | -85.9618
1.1695 | -85.9748
1.1698 | -85.9869
1.1701 | -86
1.1704 | -86.0105
1.1707 | -86.0246
1.171 | -86.0433
1.1713 | -86.0597
1.1717 | -86.0642
1.172 | -86.0729
1.1723 | -86.0822
1.1726 | -86.0971
1.1729 | -86.1103
1.1732 | -86.1202
1.1735 | -86.1226
1.1738 | -86.1326
1.1741 | -86.1389
1.1745 | -86.1438
1.1748 | -86.155
1.1751 | -86.1641
1.1754 | -86.1747
1.1757 | -86.1863
1.176 | -86.1953
1.1763 | -86.2012
1.1766 | -86.2091
1.1769 | -86.2186
1.1772 | -86.228
1.1776 | -86.2362
1.1779 | -86.2371
1.1782 | -86.2397
1.1785 | -86.2498
1.1788 | -86.2594
1.1791 | -86.2674
1.1794 | -86.2731
1.1797 | -86.2811
1.18 | -86.2924
1.1804 | -86.2964
1.1807 | -86.298
1.181 | -86.3029
1.1813 | -86.3108
1.1816 | -86.3154
1.1819 | -86.3202
1.1822 | -86.3279
1.1825 | -86.3273
1.1828 | -86.331
1.1831 | -86.337
1.1835 | -86.3439
1.1838 | -86.3468
1.1841 | -86.3511
1.1844 | -86.3525
1.1847 | -86.3559
1.185 | -86.3558
1.1853 | -86.3583
1.1856 | -86.3585
1.1859 | -86.3549
1.1863 | -86.3573
1.1866 | -86.3625
1.1869 | -86.3679
1.1872 | -86.3645
1.1875 | -86.3651
1.1878 | -86.3658
1.1881 | -86.3729
1.1884 | -86.3838
1.1887 | -86.3867
1.189 | -86.3915
1.1894 | -86.3912
1.1897 | -86.3896
1.19 | -86.39
1.1903 | -86.3862
1.1906 | -86.3866
1.1909 | -86.3849
1.1912 | -86.3834
1.1915 | -86.3821
1.1918 | -86.3795
1.1922 | -86.3772
1.1925 | -86.3806
1.1928 | -86.3833
1.1931 | -86.3828
1.1934 | -86.3814
1.1937 | -86.3815
1.194 | -86.384
1.1943 | -86.3894
1.1946 | -86.3879
1.1949 | -86.3837
1.1953 | -86.3877
1.1956 | -86.3946
1.1959 | -86.3972
1.1962 | -86.396
1.1965 | -86.3934
1.1968 | -86.387
1.1971 | -86.3798
1.1974 | -86.3749
1.1977 | -86.3681
1.1981 | -86.3587
1.1984 | -86.3513
1.1987 | -86.3492
1.199 | -86.3456
1.1993 | -86.3389
1.1996 | -86.3401
1.1999 | -86.3385
1.2002 | -86.3347
1.2005 | -86.3336
1.2008 | -86.3287
1.2012 | -86.3217
1.2015 | -86.3178
1.2018 | -86.3146
1.2021 | -86.3086
1.2024 | -86.3037
1.2027 | -86.3009
1.203 | -86.2995
1.2033 | -86.2956
1.2036 | -86.2937
1.204 | -86.2898
1.2043 | -86.2866
1.2046 | -86.2788
1.2049 | -86.2746
1.2052 | -86.2715
1.2055 | -86.271
1.2058 | -86.2707
1.2061 | -86.2679
1.2064 | -86.2655
1.2067 | -86.2602
1.2071 | -86.2563
1.2074 | -86.2535
1.2077 | -86.2542
1.208 | -86.2551
1.2083 | -86.253
1.2086 | -86.2527
1.2089 | -86.2502
1.2092 | -86.2494
1.2095 | -86.2512
1.2099 | -86.2494
1.2102 | -86.2494
1.2105 | -86.2509
1.2108 | -86.2534
1.2111 | -86.2513
1.2114 | -86.2509
1.2117 | -86.2498
1.212 | -86.2517
1.2123 | -86.2561
1.2127 | -86.2534
1.213 | -86.2557
1.2133 | -86.257
1.2136 | -86.2571
1.2139 | -86.2563
1.2142 | -86.257
1.2145 | -86.2548
1.2148 | -86.2542
1.2151 | -86.2552
1.2154 | -86.2556
1.2158 | -86.2584
1.2161 | -86.2613
1.2164 | -86.2631
1.2167 | -86.2636
1.217 | -86.266
1.2173 | -86.2728
1.2176 | -86.2773
1.2179 | -86.2791
1.2182 | -86.2798
1.2186 | -86.273
1.2189 | -86.2714
1.2192 | -86.275
1.2195 | -86.2746
1.2198 | -86.2768
1.2201 | -86.2761
1.2204 | -86.2763
1.2207 | -86.2785
1.221 | -86.2795
1.2213 | -86.2826
1.2217 | -86.2837
1.222 | -86.2884
1.2223 | -86.2914
1.2226 | -86.295
1.2229 | -86.2985
1.2232 | -86.3001
1.2235 | -86.3026
1.2238 | -86.3053
1.2241 | -86.3097
1.2245 | -86.3136
1.2248 | -86.3151
1.2251 | -86.3144
1.2254 | -86.3168
1.2257 | -86.3208
1.226 | -86.3287
1.2263 | -86.3413
1.2266 | -86.3447
1.2269 | -86.3473
1.2272 | -86.3521
1.2276 | -86.354
1.2279 | -86.3578
1.2282 | -86.3633
1.2285 | -86.3674
1.2288 | -86.3718
1.2291 | -86.3776
1.2294 | -86.3836
1.2297 | -86.3869
1.23 | -86.3884
1.2304 | -86.3968
1.2307 | -86.4002
1.231 | -86.4057
1.2313 | -86.4062
1.2316 | -86.4065
1.2319 | -86.4114
1.2322 | -86.415
1.2325 | -86.4178
1.2328 | -86.4171
1.2331 | -86.4185
1.2335 | -86.4227
1.2338 | -86.4263
1.2341 | -86.4299
1.2344 | -86.4304
1.2347 | -86.4345
1.235 | -86.4381
1.2353 | -86.4435
1.2356 | -86.4466
1.2359 | -86.4495
1.2363 | -86.4525
1.2366 | -86.4565
1.2369 | -86.4592
1.2372 | -86.4626
1.2375 | -86.4627
1.2378 | -86.4666
1.2381 | -86.4648
1.2384 | -86.47
1.2387 | -86.4776
1.239 | -86.481
1.2394 | -86.4835
1.2397 | -86.486
1.24 | -86.4819
1.2403 | -86.4851
1.2406 | -86.4857
1.2409 | -86.4866
1.2412 | -86.4893
1.2415 | -86.4905
1.2418 | -86.4918
1.2422 | -86.4908
1.2425 | -86.4938
1.2428 | -86.4987
1.2431 | -86.5
1.2434 | -86.5012
1.2437 | -86.5026
1.244 | -86.4998
1.2443 | -86.5005
1.2446 | -86.5083
1.2449 | -86.5146
1.2453 | -86.5107
1.2456 | -86.5124
1.2459 | -86.5094
1.2462 | -86.509
1.2465 | -86.5078
1.2468 | -86.5074
1.2471 | -86.5058
1.2474 | -86.5063
1.2477 | -86.4991
1.2481 | -86.492
1.2484 | -86.4878
1.2487 | -86.4849
1.249 | -86.4831
1.2493 | -86.4798
1.2496 | -86.4759
1.2499 | -86.4746
1.2502 | -86.4745
1.2505 | -86.4784
1.2508 | -86.4763
1.2512 | -86.4776
1.2515 | -86.4778
1.2518 | -86.4738
1.2521 | -86.471
1.2524 | -86.4698
1.2527 | -86.4679
1.253 | -86.4639
1.2533 | -86.4568
1.2536 | -86.451
1.254 | -86.4481
1.2543 | -86.4388
1.2546 | -86.431
1.2549 | -86.4252
1.2552 | -86.4248
1.2555 | -86.422
1.2558 | -86.4136
1.2561 | -86.4031
1.2564 | -86.4092
1.2567 | -86.4062
1.2571 | -86.4008
1.2574 | -86.3911
1.2577 | -86.3816
1.258 | -86.3768
1.2583 | -86.3744
1.2586 | -86.369
1.2589 | -86.3633
1.2592 | -86.36
1.2595 | -86.3581
1.2599 | -86.3519
1.2602 | -86.3456
1.2605 | -86.3352
1.2608 | -86.329
1.2611 | -86.3227
1.2614 | -86.3155
1.2617 | -86.3093
1.262 | -86.3026
1.2623 | -86.2946
1.2626 | -86.287
1.263 | -86.2848
1.2633 | -86.2802
1.2636 | -86.2773
1.2639 | -86.2708
1.2642 | -86.2625
1.2645 | -86.2525
1.2648 | -86.2415
1.2651 | -86.2303
1.2654 | -86.2235
1.2658 | -86.2177
1.2661 | -86.2134
1.2664 | -86.2081
1.2667 | -86.2012
1.267 | -86.1908
1.2673 | -86.1808
1.2676 | -86.175
1.2679 | -86.1669
1.2682 | -86.1602
1.2685 | -86.1543
1.2689 | -86.1481
1.2692 | -86.1369
1.2695 | -86.1283
1.2698 | -86.1211
1.2701 | -86.1155
1.2704 | -86.1085
1.2707 | -86.1048
1.271 | -86.0969
1.2713 | -86.0937
1.2717 | -86.0919
1.272 | -86.0857
1.2723 | -86.0797
1.2726 | -86.073
1.2729 | -86.0666
1.2732 | -86.0636
1.2735 | -86.0566
1.2738 | -86.0499
1.2741 | -86.0442
1.2744 | -86.0384
1.2748 | -86.0337
1.2751 | -86.0292
1.2754 | -86.0242
1.2757 | -86.0198
1.276 | -86.0105
1.2763 | -86.0063
1.2766 | -86.0006
1.2769 | -85.9962
1.2772 | -85.9907
1.2776 | -85.9865
1.2779 | -85.9816
1.2782 | -85.9781
1.2785 | -85.9758
1.2788 | -85.9721
1.2791 | -85.9671
1.2794 | -85.9617
1.2797 | -85.9526
1.28 | -85.945
1.2803 | -85.9404
1.2807 | -85.9346
1.281 | -85.9307
1.2813 | -85.9289
1.2816 | -85.928
1.2819 | -85.9255
1.2822 | -85.9249
1.2825 | -85.9235
1.2828 | -85.9225
1.2831 | -85.9194
1.2835 | -85.9121
1.2838 | -85.9094
1.2841 | -85.9059
1.2844 | -85.9045
1.2847 | -85.9029
1.285 | -85.901
1.2853 | -85.897
1.2856 | -85.8981
1.2859 | -85.8975
1.2862 | -85.8941
1.2866 | -85.8886
1.2869 | -85.8863
1.2872 | -85.8838
1.2875 | -85.8807
1.2878 | -85.8793
1.2881 | -85.8802
1.2884 | -85.8784
1.2887 | -85.8777
1.289 | -85.8784
1.2894 | -85.8778
1.2897 | -85.8736
1.29 | -85.8772
1.2903 | -85.8742
1.2906 | -85.8682
1.2909 | -85.8682
1.2912 | -85.8631
1.2915 | -85.8554
1.2918 | -85.8514
1.2921 | -85.8507
1.2925 | -85.8506
1.2928 | -85.8503
1.2931 | -85.8499
1.2934 | -85.8468
1.2937 | -85.8448
1.294 | -85.847
1.2943 | -85.8492
1.2946 | -85.8495
1.2949 | -85.851
1.2953 | -85.8516
1.2956 | -85.8512
1.2959 | -85.8505
1.2962 | -85.8509
1.2965 | -85.8503
1.2968 | -85.8528
1.2971 | -85.855
1.2974 | -85.8571
1.2977 | -85.8565
1.298 | -85.8549
1.2984 | -85.8602
1.2987 | -85.8643
1.299 | -85.8673
1.2993 | -85.8679
1.2996 | -85.8722
1.2999 | -85.8757
1.3002 | -85.8778
1.3005 | -85.8783
1.3008 | -85.8782
1.3012 | -85.8788
1.3015 | -85.8807
1.3018 | -85.8797
1.3021 | -85.8716
1.3024 | -85.8694
1.3027 | -85.8686
1.303 | -85.8696
1.3033 | -85.8717
1.3036 | -85.8728
1.3039 | -85.8746
1.3043 | -85.8755
1.3046 | -85.8774
1.3049 | -85.8794
1.3052 | -85.8802
1.3055 | -85.8799
1.3058 | -85.8777
1.3061 | -85.8796
1.3064 | -85.8879
1.3067 | -85.8978
1.3071 | -85.8993
1.3074 | -85.8974
1.3077 | -85.9009
1.308 | -85.9034
1.3083 | -85.9045
1.3086 | -85.9089
1.3089 | -85.9136
1.3092 | -85.916
1.3095 | -85.9163
1.3098 | -85.9112
1.3102 | -85.9099
1.3105 | -85.9071
1.3108 | -85.9064
1.3111 | -85.9064
1.3114 | -85.9062
1.3117 | -85.9085
1.312 | -85.9081
1.3123 | -85.9038
1.3126 | -85.8997
1.313 | -85.8989
1.3133 | -85.9013
1.3136 | -85.9041
1.3139 | -85.9064
1.3142 | -85.9054
1.3145 | -85.9042
1.3148 | -85.9039
1.3151 | -85.9021
1.3154 | -85.9027
1.3157 | -85.9012
1.3161 | -85.8997
1.3164 | -85.9005
1.3167 | -85.8995
1.317 | -85.896
1.3173 | -85.8949
1.3176 | -85.8921
1.3179 | -85.8874
1.3182 | -85.8825
1.3185 | -85.8773
1.3189 | -85.8767
1.3192 | -85.8763
1.3195 | -85.8745
1.3198 | -85.8726
1.3201 | -85.8702
1.3204 | -85.8727
1.3207 | -85.8747
1.321 | -85.8698
1.3213 | -85.8667
1.3216 | -85.865
1.322 | -85.8678
1.3223 | -85.8684
1.3226 | -85.8644
1.3229 | -85.8602
1.3232 | -85.8584
1.3235 | -85.8582
1.3238 | -85.8535
1.3241 | -85.8496
1.3244 | -85.8458
1.3248 | -85.8421
1.3251 | -85.8387
1.3254 | -85.8345
1.3257 | -85.8297
1.326 | -85.8227
1.3263 | -85.8186
1.3266 | -85.8155
1.3269 | -85.8073
1.3272 | -85.8015
1.3275 | -85.7988
1.3279 | -85.7987
1.3282 | -85.7955
1.3285 | -85.7918
1.3288 | -85.7884
1.3291 | -85.7838
1.3294 | -85.7789
1.3297 | -85.7794
1.33 | -85.7742
1.3303 | -85.77
1.3307 | -85.7657
1.331 | -85.7608
1.3313 | -85.7554
1.3316 | -85.7486
1.3319 | -85.7461
1.3322 | -85.7434
1.3325 | -85.7415
1.3328 | -85.7375
1.3331 | -85.7329
1.3334 | -85.7263
1.3338 | -85.7249
1.3341 | -85.7209
1.3344 | -85.7174
1.3347 | -85.7076
1.335 | -85.7061
1.3353 | -85.7048
1.3356 | -85.7005
1.3359 | -85.6948
1.3362 | -85.6939
1.3366 | -85.6922
1.3369 | -85.6896
1.3372 | -85.6854
1.3375 | -85.6833
1.3378 | -85.6803
1.3381 | -85.6819
1.3384 | -85.6837
1.3387 | -85.6803
1.339 | -85.674
1.3393 | -85.6717
1.3397 | -85.6702
1.34 | -85.6695
1.3403 | -85.6669
1.3406 | -85.6642
1.3409 | -85.6623
1.3412 | -85.6606
1.3415 | -85.6582
1.3418 | -85.6542
1.3421 | -85.6531
1.3425 | -85.6521
1.3428 | -85.6498
1.3431 | -85.649
1.3434 | -85.6495
1.3437 | -85.6512
1.344 | -85.6504
1.3443 | -85.6496
1.3446 | -85.649
1.3449 | -85.6509
1.3453 | -85.6542
1.3456 | -85.6576
1.3459 | -85.659
1.3462 | -85.6599
1.3465 | -85.6611
1.3468 | -85.6659
1.3471 | -85.6759
1.3474 | -85.6773
1.3477 | -85.6798
1.348 | -85.6824
1.3484 | -85.6875
1.3487 | -85.6922
1.349 | -85.6944
1.3493 | -85.6968
1.3496 | -85.6988
1.3499 | -85.7011
1.3502 | -85.7042
1.3505 | -85.7062
1.3508 | -85.7073
1.3512 | -85.7124
1.3515 | -85.7198
1.3518 | -85.7251
1.3521 | -85.7299
1.3524 | -85.7369
1.3527 | -85.7447
1.353 | -85.7493
1.3533 | -85.7539
1.3536 | -85.7601
1.3539 | -85.7658
1.3543 | -85.7715
1.3546 | -85.7769
1.3549 | -85.7828
1.3552 | -85.7894
1.3555 | -85.7952
1.3558 | -85.8027
1.3561 | -85.812
1.3564 | -85.8234
1.3567 | -85.8341
1.3571 | -85.8391
1.3574 | -85.8449
1.3577 | -85.8477
1.358 | -85.8532
1.3583 | -85.8597
1.3586 | -85.87
1.3589 | -85.879
1.3592 | -85.8849
1.3595 | -85.8894
1.3598 | -85.8906
1.3602 | -85.8958
1.3605 | -85.9031
1.3608 | -85.9132
1.3611 | -85.9227
1.3614 | -85.9305
1.3617 | -85.9427
1.362 | -85.951
1.3623 | -85.9604
1.3626 | -85.9669
1.363 | -85.9778
1.3633 | -85.9869
1.3636 | -85.9966
1.3639 | -86.0058
1.3642 | -86.0165
1.3645 | -86.0263
1.3648 | -86.0344
1.3651 | -86.0446
1.3654 | -86.0545
1.3657 | -86.0654
1.3661 | -86.0747
1.3664 | -86.083
1.3667 | -86.0936
1.367 | -86.1044
1.3673 | -86.1154
1.3676 | -86.1273
1.3679 | -86.1376
1.3682 | -86.1518
1.3685 | -86.1639
1.3689 | -86.1723
1.3692 | -86.1782
1.3695 | -86.1887
1.3698 | -86.1983
1.3701 | -86.2084
1.3704 | -86.2185
1.3707 | -86.2305
1.371 | -86.2417
1.3713 | -86.2518
1.3716 | -86.2606
1.372 | -86.2714
1.3723 | -86.2873
1.3726 | -86.3013
1.3729 | -86.3095
1.3732 | -86.3185
1.3735 | -86.3257
1.3738 | -86.3327
1.3741 | -86.3428
1.3744 | -86.3521
1.3748 | -86.3622
1.3751 | -86.3696
1.3754 | -86.3778
1.3757 | -86.3877
1.376 | -86.3961
1.3763 | -86.4073
1.3766 | -86.4169
1.3769 | -86.427
1.3772 | -86.4389
1.3775 | -86.4503
1.3779 | -86.4594
1.3782 | -86.4674
1.3785 | -86.473
1.3788 | -86.4808
1.3791 | -86.4893
1.3794 | -86.5004
1.3797 | -86.5027
1.38 | -86.5034
1.3803 | -86.5053
1.3807 | -86.5075
1.381 | -86.5103
1.3813 | -86.5168
1.3816 | -86.5232
1.3819 | -86.5292
1.3822 | -86.5358
1.3825 | -86.542
1.3828 | -86.5522
1.3831 | -86.5597
1.3834 | -86.5648
1.3838 | -86.5709
1.3841 | -86.5827
1.3844 | -86.5883
1.3847 | -86.5994
1.385 | -86.6101
1.3853 | -86.6227
1.3856 | -86.6251
1.3859 | -86.6268
1.3862 | -86.6296
1.3866 | -86.6321
1.3869 | -86.6333
1.3872 | -86.6304
1.3875 | -86.6326
1.3878 | -86.6345
1.3881 | -86.6357
1.3884 | -86.6363
1.3887 | -86.6355
1.389 | -86.6373
1.3893 | -86.636
1.3897 | -86.6362
1.39 | -86.6396
1.3903 | -86.642
1.3906 | -86.643
1.3909 | -86.6404
1.3912 | -86.6357
1.3915 | -86.6376
1.3918 | -86.6386
1.3921 | -86.6351
1.3925 | -86.6347
1.3928 | -86.6344
1.3931 | -86.6362
1.3934 | -86.6369
1.3937 | -86.6349
1.394 | -86.6338
1.3943 | -86.631
1.3946 | -86.6299
1.3949 | -86.6322
1.3952 | -86.6287
1.3956 | -86.6234
1.3959 | -86.6235
1.3962 | -86.6333
1.3965 | -86.6365
1.3968 | -86.6321
1.3971 | -86.6265
1.3974 | -86.6192
1.3977 | -86.6147
1.398 | -86.6121
1.3984 | -86.6076
1.3987 | -86.6051
1.399 | -86.602
1.3993 | -86.5974
1.3996 | -86.5945
1.3999 | -86.5902
1.4002 | -86.5901
1.4005 | -86.5862
1.4008 | -86.5803
1.4011 | -86.5733
1.4015 | -86.5718
1.4018 | -86.5695
1.4021 | -86.5647
1.4024 | -86.559
1.4027 | -86.5545
1.403 | -86.5465
1.4033 | -86.5343
1.4036 | -86.527
1.4039 | -86.5201
1.4043 | -86.5119
1.4046 | -86.5101
1.4049 | -86.509
1.4052 | -86.507
1.4055 | -86.5031
1.4058 | -86.4969
1.4061 | -86.4901
1.4064 | -86.484
1.4067 | -86.4779
1.407 | -86.4747
1.4074 | -86.474
1.4077 | -86.4707
1.408 | -86.4606
1.4083 | -86.4504
1.4086 | -86.4416
1.4089 | -86.4326
1.4092 | -86.4258
1.4095 | -86.4192
1.4098 | -86.412
1.4102 | -86.4059
1.4105 | -86.398
1.4108 | -86.3921
1.4111 | -86.3876
1.4114 | -86.3807
1.4117 | -86.3741
1.412 | -86.3685
1.4123 | -86.3599
1.4126 | -86.3535
1.4129 | -86.3483
1.4133 | -86.3412
1.4136 | -86.3367
1.4139 | -86.3327
1.4142 | -86.3246
1.4145 | -86.3191
1.4148 | -86.3076
1.4151 | -86.2938
1.4154 | -86.2898
1.4157 | -86.2845
1.4161 | -86.278
1.4164 | -86.2719
1.4167 | -86.2681
1.417 | -86.2639
1.4173 | -86.2601
1.4176 | -86.2542
1.4179 | -86.2486
1.4182 | -86.244
1.4185 | -86.2392
1.4188 | -86.2342
1.4192 | -86.227
1.4195 | -86.2233
1.4198 | -86.2177
1.4201 | -86.2112
1.4204 | -86.2032
1.4207 | -86.1983
1.421 | -86.195
1.4213 | -86.1915
1.4216 | -86.1885
1.422 | -86.1843
1.4223 | -86.1735
1.4226 | -86.1582
1.4229 | -86.152
1.4232 | -86.1479
1.4235 | -86.1416
1.4238 | -86.136
1.4241 | -86.1306
1.4244 | -86.1296
1.4247 | -86.1256
1.4251 | -86.1199
1.4254 | -86.1159
1.4257 | -86.1121
1.426 | -86.0993
1.4263 | -86.0911
1.4266 | -86.0877
1.4269 | -86.0834
1.4272 | -86.0779
1.4275 | -86.075
1.4279 | -86.0733
1.4282 | -86.0697
1.4285 | -86.0659
1.4288 | -86.0618
1.4291 | -86.0581
1.4294 | -86.055
1.4297 | -86.0543
1.43 | -86.0492
1.4303 | -86.0444
1.4306 | -86.0397
1.431 | -86.043
1.4313 | -86.0407
1.4316 | -86.0382
1.4319 | -86.0345
1.4322 | -86.0313
1.4325 | -86.0299
1.4328 | -86.0267
1.4331 | -86.0225
1.4334 | -86.0224
1.4338 | -86.0202
1.4341 | -86.0184
1.4344 | -86.0143
1.4347 | -86.01
1.435 | -86.0094
1.4353 | -86.0077
1.4356 | -86.0056
1.4359 | -86.0081
1.4362 | -86.0069
1.4365 | -86.006
1.4369 | -86.0044
1.4372 | -86.004
1.4375 | -86.0029
1.4378 | -86.0008
1.4381 | -85.9992
1.4384 | -85.9978
1.4387 | -85.9971
1.439 | -85.997
1.4393 | -85.9963
1.4397 | -85.9943
1.44 | -85.9931
1.4403 | -85.9916
1.4406 | -85.9921
1.4409 | -85.9932
1.4412 | -85.9962
1.4415 | -85.9983
1.4418 | -85.9968
1.4421 | -85.9956
1.4424 | -85.9963
1.4428 | -85.9929
1.4431 | -85.9936
1.4434 | -85.9935
1.4437 | -85.9933
1.444 | -85.9936
1.4443 | -85.9934
1.4446 | -85.9919
1.4449 | -85.9906
1.4452 | -85.9907
1.4456 | -85.9894
1.4459 | -85.9895
1.4462 | -85.9912
1.4465 | -85.993
1.4468 | -85.9952
1.4471 | -85.9974
1.4474 | -85.9991
1.4477 | -86.0008
1.448 | -86.0034
1.4483 | -86.0059
1.4487 | -86.0069
1.449 | -86.0032
1.4493 | -85.9992
1.4496 | -85.9958
1.4499 | -85.989
1.4502 | -85.9892
1.4505 | -85.9925
1.4508 | -85.99
1.4511 | -85.9905
1.4515 | -85.994
1.4518 | -85.9952
1.4521 | -85.9969
1.4524 | -85.9981
1.4527 | -85.9949
1.453 | -85.9904
1.4533 | -85.9872
1.4536 | -85.9873
1.4539 | -85.9876
1.4542 | -85.9849
1.4546 | -85.9858
1.4549 | -85.9864
1.4552 | -85.9891
1.4555 | -85.993
1.4558 | -85.9971
1.4561 | -85.9972
1.4564 | -85.9982
1.4567 | -85.9989
1.457 | -86.0009
1.4574 | -86.0034
1.4577 | -86.0088
1.458 | -86.0165
1.4583 | -86.0193
1.4586 | -86.0206
1.4589 | -86.0229
1.4592 | -86.0236
1.4595 | -86.0211
1.4598 | -86.0217
1.4601 | -86.0231
1.4605 | -86.0233
1.4608 | -86.0243
1.4611 | -86.0254
1.4614 | -86.0276
1.4617 | -86.0281
1.462 | -86.0287
1.4623 | -86.0281
1.4626 | -86.0269
1.4629 | -86.0227
1.4633 | -86.018
1.4636 | -86.0179
1.4639 | -86.0187
1.4642 | -86.0187
1.4645 | -86.0192
1.4648 | -86.0201
1.4651 | -86.0207
1.4654 | -86.0201
1.4657 | -86.0193
1.466 | -86.0177
1.4664 | -86.0178
1.4667 | -86.0188
1.467 | -86.0197
1.4673 | -86.0208
1.4676 | -86.0237
1.4679 | -86.0276
1.4682 | -86.0287
1.4685 | -86.0303
1.4688 | -86.0299
1.4692 | -86.0322
1.4695 | -86.0335
1.4698 | -86.0341
1.4701 | -86.0338
1.4704 | -86.0336
1.4707 | -86.0342
1.471 | -86.0382
1.4713 | -86.0384
1.4716 | -86.037
1.4719 | -86.0381
1.4723 | -86.0415
1.4726 | -86.0462
1.4729 | -86.0525
1.4732 | -86.0545
1.4735 | -86.051
1.4738 | -86.051
1.4741 | -86.0508
1.4744 | -86.0524
1.4747 | -86.0536
1.4751 | -86.0532
1.4754 | -86.0544
1.4757 | -86.0555
1.476 | -86.0579
1.4763 | -86.0564
1.4766 | -86.0576
1.4769 | -86.0577
1.4772 | -86.0593
1.4775 | -86.0598
1.4778 | -86.0618
1.4782 | -86.0619
1.4785 | -86.0632
1.4788 | -86.0639
1.4791 | -86.0642
1.4794 | -86.0639
1.4797 | -86.0643
1.48 | -86.0682
1.4803 | -86.07
1.4806 | -86.0733
1.481 | -86.0761
1.4813 | -86.076
1.4816 | -86.0758
1.4819 | -86.0747
1.4822 | -86.0731
1.4825 | -86.0727
1.4828 | -86.0724
1.4831 | -86.0704
1.4834 | -86.0707
1.4837 | -86.0723
1.4841 | -86.0717
1.4844 | -86.0718
1.4847 | -86.0709
1.485 | -86.0725
1.4853 | -86.0727
1.4856 | -86.0747
1.4859 | -86.0733
1.4862 | -86.0733
1.4865 | -86.0751
1.4869 | -86.0804
1.4872 | -86.0832
1.4875 | -86.0883
1.4878 | -86.0848
1.4881 | -86.084
1.4884 | -86.0825
1.4887 | -86.0796
1.489 | -86.0782
1.4893 | -86.0752
1.4896 | -86.071
1.49 | -86.0717
1.4903 | -86.0739
1.4906 | -86.0739
1.4909 | -86.074
1.4912 | -86.0739
1.4915 | -86.0759
1.4918 | -86.0788
1.4921 | -86.0791
1.4924 | -86.0795
1.4928 | -86.0827
1.4931 | -86.0793
1.4934 | -86.074
1.4937 | -86.0726
1.494 | -86.069
1.4943 | -86.0699
1.4946 | -86.0687
1.4949 | -86.0697
1.4952 | -86.0727
1.4955 | -86.0707
1.4959 | -86.0661
1.4962 | -86.0643
1.4965 | -86.0667
1.4968 | -86.0683
1.4971 | -86.0625
1.4974 | -86.0591
1.4977 | -86.062
1.498 | -86.0636
1.4983 | -86.0634
1.4987 | -86.0641
1.499 | -86.0632
1.4993 | -86.0619
1.4996 | -86.0628
1.4999 | -86.0617
1.5002 | -86.0571
1.5005 | -86.0572
1.5008 | -86.0567
1.5011 | -86.0556
1.5014 | -86.0545
1.5018 | -86.0521
1.5021 | -86.0464
1.5024 | -86.0465
1.5027 | -86.0457
1.503 | -86.046
1.5033 | -86.0451
1.5036 | -86.0457
1.5039 | -86.0451
1.5042 | -86.0478
1.5046 | -86.0526
1.5049 | -86.0514
1.5052 | -86.0495
1.5055 | -86.0476
1.5058 | -86.0496
1.5061 | -86.0467
1.5064 | -86.0421
1.5067 | -86.0389
1.507 | -86.0357
1.5073 | -86.0359
1.5077 | -86.0355
1.508 | -86.0336
1.5083 | -86.0329
1.5086 | -86.0304
1.5089 | -86.0279
1.5092 | -86.027
1.5095 | -86.0272
1.5098 | -86.0289
1.5101 | -86.0296
1.5105 | -86.0311
1.5108 | -86.0317
1.5111 | -86.0279
1.5114 | -86.0227
1.5117 | -86.0165
1.512 | -86.0143
1.5123 | -86.0145
1.5126 | -86.013
1.5129 | -86.0117
1.5133 | -86.0101
1.5136 | -86.0072
1.5139 | -86.0055
1.5142 | -86.0047
1.5145 | -85.9993
1.5148 | -85.993
1.5151 | -85.9912
1.5154 | -85.9918
1.5157 | -85.9909
1.516 | -85.9882
1.5164 | -85.9858
1.5167 | -85.9839
1.517 | -85.9856
1.5173 | -85.9876
1.5176 | -85.9843
1.5179 | -85.9819
1.5182 | -85.9787
1.5185 | -85.9751
1.5188 | -85.9725
1.5192 | -85.9698
1.5195 | -85.9652
1.5198 | -85.9619
1.5201 | -85.9581
1.5204 | -85.9534
1.5207 | -85.9509
1.521 | -85.9487
1.5213 | -85.9456
1.5216 | -85.9432
1.5219 | -85.9408
1.5223 | -85.9377
1.5226 | -85.9349
1.5229 | -85.931
1.5232 | -85.9307
1.5235 | -85.9267
1.5238 | -85.9199
1.5241 | -85.9134
1.5244 | -85.9097
1.5247 | -85.9077
1.5251 | -85.9106
1.5254 | -85.9161
1.5257 | -85.9178
1.526 | -85.9168
1.5263 | -85.9143
1.5266 | -85.9089
1.5269 | -85.9045
1.5272 | -85.9018
1.5275 | -85.8984
1.5278 | -85.8944
1.5282 | -85.8923
1.5285 | -85.8893
1.5288 | -85.887
1.5291 | -85.8913
1.5294 | -85.8885
1.5297 | -85.8822
1.53 | -85.8788
1.5303 | -85.8759
1.5306 | -85.8716
1.531 | -85.8679
1.5313 | -85.8642
1.5316 | -85.8591
1.5319 | -85.8568
1.5322 | -85.8527
1.5325 | -85.848
1.5328 | -85.8444
1.5331 | -85.841
1.5334 | -85.8386
1.5337 | -85.8336
1.5341 | -85.8286
1.5344 | -85.8209
1.5347 | -85.8173
1.535 | -85.8142
1.5353 | -85.8103
1.5356 | -85.8067
1.5359 | -85.8035
1.5362 | -85.8007
1.5365 | -85.7995
1.5369 | -85.7968
1.5372 | -85.7928
1.5375 | -85.7887
1.5378 | -85.7779
1.5381 | -85.7745
1.5384 | -85.7732
1.5387 | -85.7698
1.539 | -85.7658
1.5393 | -85.7646
1.5396 | -85.7598
1.54 | -85.7517
1.5403 | -85.7468
1.5406 | -85.7438
1.5409 | -85.7442
1.5412 | -85.7387
1.5415 | -85.7335
1.5418 | -85.731
1.5421 | -85.7287
1.5424 | -85.7268
1.5428 | -85.7256
1.5431 | -85.7219
1.5434 | -85.7199
1.5437 | -85.7197
1.544 | -85.7146
1.5443 | -85.7116
1.5446 | -85.7086
1.5449 | -85.7081
1.5452 | -85.7062
1.5455 | -85.7014
1.5459 | -85.698
1.5462 | -85.6934
1.5465 | -85.6882
1.5468 | -85.6866
1.5471 | -85.6871
1.5474 | -85.6831
1.5477 | -85.6816
1.548 | -85.6797
1.5483 | -85.6743
1.5487 | -85.6688
1.549 | -85.6678
1.5493 | -85.6667
1.5496 | -85.6626
1.5499 | -85.664
1.5502 | -85.6619
1.5505 | -85.6595
1.5508 | -85.6551
1.5511 | -85.6523
1.5514 | -85.6503
1.5518 | -85.6491
1.5521 | -85.6474
1.5524 | -85.6462
1.5527 | -85.6452
1.553 | -85.6457
1.5533 | -85.6454
1.5536 | -85.6487
1.5539 | -85.6519
1.5542 | -85.6541
1.5546 | -85.6605
1.5549 | -85.6614
1.5552 | -85.6606
1.5555 | -85.6636
1.5558 | -85.6632
1.5561 | -85.6615
1.5564 | -85.6602
1.5567 | -85.6575
1.557 | -85.6513
1.5573 | -85.6433
1.5577 | -85.6447
1.558 | -85.6489
1.5583 | -85.6493
1.5586 | -85.6476
1.5589 | -85.6493
1.5592 | -85.6487
1.5595 | -85.6506
1.5598 | -85.6469
1.5601 | -85.6424
1.5605 | -85.6424
1.5608 | -85.642
1.5611 | -85.6436
1.5614 | -85.6445
1.5617 | -85.6417
1.562 | -85.6365
1.5623 | -85.6316
1.5626 | -85.6287
1.5629 | -85.628
1.5632 | -85.6279
1.5636 | -85.6274
1.5639 | -85.6262
1.5642 | -85.6262
1.5645 | -85.6277
1.5648 | -85.6302
1.5651 | -85.6299
1.5654 | -85.6295
1.5657 | -85.6257
1.566 | -85.6255
1.5664 | -85.6252
1.5667 | -85.6241
1.567 | -85.6224
1.5673 | -85.6197
1.5676 | -85.619
1.5679 | -85.6196
1.5682 | -85.6196
1.5685 | -85.6183
1.5688 | -85.6204
1.5691 | -85.6248
1.5695 | -85.6228
1.5698 | -85.6213
1.5701 | -85.6211
1.5704 | -85.6161
1.5707 | -85.6142
1.571 | -85.6147
1.5713 | -85.6147
1.5716 | -85.6147
1.5719 | -85.6126
1.5723 | -85.6109
1.5726 | -85.6102
1.5729 | -85.6104
1.5732 | -85.6093
1.5735 | -85.6089
1.5738 | -85.6091
1.5741 | -85.6075
1.5744 | -85.6058
1.5747 | -85.6045
1.575 | -85.6
1.5754 | -85.5979
1.5757 | -85.5971
1.576 | -85.5965
1.5763 | -85.5967
1.5766 | -85.5939
1.5769 | -85.5813
1.5772 | -85.5801
1.5775 | -85.5779
1.5778 | -85.5788
1.5782 | -85.5791
1.5785 | -85.5769
1.5788 | -85.5746
1.5791 | -85.5718
1.5794 | -85.5676
1.5797 | -85.5679
1.58 | -85.5696
1.5803 | -85.5676
1.5806 | -85.5655
1.5809 | -85.5647
1.5813 | -85.5629
1.5816 | -85.5618
1.5819 | -85.561
1.5822 | -85.5621
1.5825 | -85.5606
1.5828 | -85.5595
1.5831 | -85.5567
1.5834 | -85.5532
1.5837 | -85.548
1.5841 | -85.5466
1.5844 | -85.5453
1.5847 | -85.5441
1.585 | -85.5429
1.5853 | -85.5407
1.5856 | -85.5379
1.5859 | -85.533
1.5862 | -85.5267
1.5865 | -85.5245
1.5868 | -85.5228
1.5872 | -85.5211
1.5875 | -85.5215
1.5878 | -85.5181
1.5881 | -85.5147
1.5884 | -85.5123
1.5887 | -85.5091
1.589 | -85.5008
1.5893 | -85.491
1.5896 | -85.4854
1.59 | -85.4805
1.5903 | -85.4755
1.5906 | -85.4711
1.5909 | -85.4671
1.5912 | -85.4633
1.5915 | -85.4592
1.5918 | -85.4543
1.5921 | -85.4509
1.5924 | -85.4465
1.5927 | -85.4419
1.5931 | -85.4371
1.5934 | -85.4328
1.5937 | -85.4278
1.594 | -85.4201
1.5943 | -85.4146
1.5946 | -85.4085
1.5949 | -85.4033
1.5952 | -85.4007
1.5955 | -85.4008
1.5959 | -85.3967
1.5962 | -85.3952
1.5965 | -85.3943
1.5968 | -85.3915
1.5971 | -85.3902
1.5974 | -85.3886
1.5977 | -85.3872
1.598 | -85.3823
1.5983 | -85.3769
1.5986 | -85.3715
1.599 | -85.3666
1.5993 | -85.3611
1.5996 | -85.3542
1.5999 | -85.3471
1.6002 | -85.3445
1.6005 | -85.3404
1.6008 | -85.3381
1.6011 | -85.3374
1.6014 | -85.3343
1.6018 | -85.3331
1.6021 | -85.3282
1.6024 | -85.325
1.6027 | -85.3211
1.603 | -85.3168
1.6033 | -85.3143
1.6036 | -85.3134
1.6039 | -85.315
1.6042 | -85.3137
1.6045 | -85.3097
1.6049 | -85.306
1.6052 | -85.3021
1.6055 | -85.3026
1.6058 | -85.2975
1.6061 | -85.2901
1.6064 | -85.2851
1.6067 | -85.2802
1.607 | -85.2748
1.6073 | -85.2698
1.6077 | -85.2656
1.608 | -85.2601
1.6083 | -85.2532
1.6086 | -85.2528
1.6089 | -85.2492
1.6092 | -85.2451
1.6095 | -85.2416
1.6098 | -85.2374
1.6101 | -85.2335
1.6104 | -85.2338
1.6108 | -85.2314
1.6111 | -85.2238
1.6114 | -85.2151
1.6117 | -85.2104
1.612 | -85.2064
1.6123 | -85.2026
1.6126 | -85.2014
1.6129 | -85.1948
1.6132 | -85.1895
1.6136 | -85.1867
1.6139 | -85.184
1.6142 | -85.1824
1.6145 | -85.1793
1.6148 | -85.1756
1.6151 | -85.1759
1.6154 | -85.176
1.6157 | -85.1737
1.616 | -85.1722
1.6163 | -85.1674
1.6167 | -85.1632
1.617 | -85.1597
1.6173 | -85.1552
1.6176 | -85.1507
1.6179 | -85.1471
1.6182 | -85.1422
1.6185 | -85.1394
1.6188 | -85.1372
1.6191 | -85.1345
1.6195 | -85.1307
1.6198 | -85.1286
1.6201 | -85.1283
1.6204 | -85.1265
1.6207 | -85.1216
1.621 | -85.1185
1.6213 | -85.1162
1.6216 | -85.1118
1.6219 | -85.1072
1.6222 | -85.0997
1.6226 | -85.0885
1.6229 | -85.0821
1.6232 | -85.0776
1.6235 | -85.0708
1.6238 | -85.0652
1.6241 | -85.067
1.6244 | -85.0663
1.6247 | -85.0633
1.625 | -85.0603
1.6254 | -85.0558
1.6257 | -85.0515
1.626 | -85.0481
1.6263 | -85.0454
1.6266 | -85.0428
1.6269 | -85.0414
1.6272 | -85.0386
1.6275 | -85.0346
1.6278 | -85.0322
1.6281 | -85.0303
1.6285 | -85.0276
1.6288 | -85.0243
1.6291 | -85.0214
1.6294 | -85.0189
1.6297 | -85.0159
1.63 | -85.0135
1.6303 | -85.01
1.6306 | -85.0071
1.6309 | -85.0019
1.6313 | -85
1.6316 | -84.9969
1.6319 | -84.9924
1.6322 | -84.9909
1.6325 | -84.9894
1.6328 | -84.9872
1.6331 | -84.986
1.6334 | -84.9837
1.6337 | -84.9821
1.6341 | -84.9831
1.6344 | -84.9835
1.6347 | -84.9808
1.635 | -84.9763
1.6353 | -84.9742
1.6356 | -84.9709
1.6359 | -84.9644
1.6362 | -84.9617
1.6365 | -84.9583
1.6368 | -84.9579
1.6372 | -84.9563
1.6375 | -84.9541
1.6378 | -84.9475
1.6381 | -84.9395
1.6384 | -84.9372
1.6387 | -84.9337
1.639 | -84.9316
1.6393 | -84.9281
1.6396 | -84.9262
1.64 | -84.9263
1.6403 | -84.9257
1.6406 | -84.9238
1.6409 | -84.9224
1.6412 | -84.9211
1.6415 | -84.9192
1.6418 | -84.9153
1.6421 | -84.912
1.6424 | -84.9109
1.6427 | -84.9089
1.6431 | -84.9054
1.6434 | -84.9016
1.6437 | -84.8998
1.644 | -84.8985
1.6443 | -84.8975
1.6446 | -84.8981
1.6449 | -84.8987
1.6452 | -84.9001
1.6455 | -84.8984
1.6459 | -84.8966
1.6462 | -84.8955
1.6465 | -84.895
1.6468 | -84.8915
1.6471 | -84.8885
1.6474 | -84.885
1.6477 | -84.884
1.648 | -84.8821
1.6483 | -84.8829
1.6486 | -84.881
1.649 | -84.885
1.6493 | -84.8841
1.6496 | -84.8806
1.6499 | -84.8766
1.6502 | -84.8738
1.6505 | -84.8719
1.6508 | -84.8694
1.6511 | -84.8704
1.6514 | -84.8689
1.6518 | -84.8674
1.6521 | -84.865
1.6524 | -84.8636
1.6527 | -84.864
1.653 | -84.8637
1.6533 | -84.8595
1.6536 | -84.8554
1.6539 | -84.8545
1.6542 | -84.8531
1.6545 | -84.8525
1.6549 | -84.8511
1.6552 | -84.85
1.6555 | -84.8513
1.6558 | -84.8513
1.6561 | -84.85
1.6564 | -84.8484
1.6567 | -84.8464
1.657 | -84.8456
1.6573 | -84.8443
1.6577 | -84.8451
1.658 | -84.8453
1.6583 | -84.8442
1.6586 | -84.843
1.6589 | -84.8405
1.6592 | -84.8408
1.6595 | -84.8386
1.6598 | -84.8367
1.6601 | -84.8388
1.6604 | -84.8409
1.6608 | -84.8405
1.6611 | -84.8404
1.6614 | -84.841
1.6617 | -84.8403
1.662 | -84.8399
1.6623 | -84.8408
1.6626 | -84.841
1.6629 | -84.8409
1.6632 | -84.8338
1.6636 | -84.8289
1.6639 | -84.8273
1.6642 | -84.8272
1.6645 | -84.8262
1.6648 | -84.8247
1.6651 | -84.8238
1.6654 | -84.8235
1.6657 | -84.8221
1.666 | -84.8206
1.6663 | -84.8228
1.6667 | -84.8226
1.667 | -84.8204
1.6673 | -84.8201
1.6676 | -84.8199
1.6679 | -84.8201
1.6682 | -84.821
1.6685 | -84.8215
1.6688 | -84.8225
1.6691 | -84.8248
1.6695 | -84.8299
1.6698 | -84.834
1.6701 | -84.8353
1.6704 | -84.8356
1.6707 | -84.8379
1.671 | -84.839
1.6713 | -84.8395
1.6716 | -84.8395
1.6719 | -84.8373
1.6722 | -84.8377
1.6726 | -84.8395
1.6729 | -84.8415
1.6732 | -84.8424
1.6735 | -84.8429
1.6738 | -84.8414
1.6741 | -84.8422
1.6744 | -84.8453
1.6747 | -84.8499
1.675 | -84.8531
1.6754 | -84.8544
1.6757 | -84.8571
1.676 | -84.8569
1.6763 | -84.8579
1.6766 | -84.8583
1.6769 | -84.8605
1.6772 | -84.8619
1.6775 | -84.8635
1.6778 | -84.8642
1.6781 | -84.8648
1.6785 | -84.8647
1.6788 | -84.8689
1.6791 | -84.87
1.6794 | -84.8695
1.6797 | -84.871
1.68 | -84.8747
1.6803 | -84.8785
1.6806 | -84.8784
1.6809 | -84.8803
1.6813 | -84.882
1.6816 | -84.8845
1.6819 | -84.889
1.6822 | -84.8923
1.6825 | -84.8952
1.6828 | -84.8993
1.6831 | -84.9014
1.6834 | -84.9041
1.6837 | -84.9097
1.684 | -84.9132
1.6844 | -84.9182
1.6847 | -84.9202
1.685 | -84.923
1.6853 | -84.9263
1.6856 | -84.9318
1.6859 | -84.9357
1.6862 | -84.936
1.6865 | -84.9353
1.6868 | -84.9334
1.6872 | -84.933
1.6875 | -84.9299
1.6878 | -84.9329
1.6881 | -84.9365
1.6884 | -84.9412
1.6887 | -84.9465
1.689 | -84.9482
1.6893 | -84.9501
1.6896 | -84.9491
1.6899 | -84.9471
1.6903 | -84.9493
1.6906 | -84.9585
1.6909 | -84.964
1.6912 | -84.9671
1.6915 | -84.9705
1.6918 | -84.9732
1.6921 | -84.9734
1.6924 | -84.9722
1.6927 | -84.974
1.6931 | -84.9758
1.6934 | -84.9791
1.6937 | -84.9803
1.694 | -84.9797
1.6943 | -84.979
1.6946 | -84.9798
1.6949 | -84.9792
1.6952 | -84.9778
1.6955 | -84.9768
1.6958 | -84.9687
1.6962 | -84.9614
1.6965 | -84.9613
1.6968 | -84.9616
1.6971 | -84.9609
1.6974 | -84.9629
1.6977 | -84.9613
1.698 | -84.9621
1.6983 | -84.963
1.6986 | -84.9635
1.699 | -84.9647
1.6993 | -84.9653
1.6996 | -84.9653
1.6999 | -84.9672
1.7002 | -84.9619
1.7005 | -84.9571
1.7008 | -84.9556
1.7011 | -84.9553
1.7014 | -84.9557
1.7017 | -84.9552
1.7021 | -84.953
1.7024 | -84.9508
1.7027 | -84.9498
1.703 | -84.9484
1.7033 | -84.9432
1.7036 | -84.9406
1.7039 | -84.9454
1.7042 | -84.9513
1.7045 | -84.9527
1.7049 | -84.9528
1.7052 | -84.9544
1.7055 | -84.9535
1.7058 | -84.9523
1.7061 | -84.9515
1.7064 | -84.9505
1.7067 | -84.9512
1.707 | -84.9497
1.7073 | -84.9488
1.7076 | -84.9484
1.708 | -84.9496
1.7083 | -84.9483
1.7086 | -84.9473
1.7089 | -84.9481
1.7092 | -84.9473
1.7095 | -84.9458
1.7098 | -84.9445
1.7101 | -84.9425
1.7104 | -84.9402
1.7108 | -84.9376
1.7111 | -84.9328
1.7114 | -84.9312
1.7117 | -84.9282
1.712 | -84.9259
1.7123 | -84.9196
1.7126 | -84.9103
1.7129 | -84.9055
1.7132 | -84.9023
1.7135 | -84.8994
1.7139 | -84.8972
1.7142 | -84.8944
1.7145 | -84.8902
1.7148 | -84.8877
1.7151 | -84.885
1.7154 | -84.8832
1.7157 | -84.8783
1.716 | -84.8743
1.7163 | -84.8706
1.7167 | -84.8662
1.717 | -84.8623
1.7173 | -84.8579
1.7176 | -84.854
1.7179 | -84.8493
1.7182 | -84.8427
1.7185 | -84.8376
1.7188 | -84.8318
1.7191 | -84.8261
1.7194 | -84.819
1.7198 | -84.8135
1.7201 | -84.8048
1.7204 | -84.8001
1.7207 | -84.7946
1.721 | -84.7895
1.7213 | -84.7846
1.7216 | -84.7793
1.7219 | -84.7752
1.7222 | -84.7713
1.7226 | -84.7669
1.7229 | -84.7616
1.7232 | -84.7576
1.7235 | -84.7522
1.7238 | -84.7465
1.7241 | -84.7423
1.7244 | -84.7395
1.7247 | -84.7331
1.725 | -84.726
1.7253 | -84.7201
1.7257 | -84.7158
1.726 | -84.7106
1.7263 | -84.7052
1.7266 | -84.6985
1.7269 | -84.6889
1.7272 | -84.6835
1.7275 | -84.6793
1.7278 | -84.6743
1.7281 | -84.668
1.7285 | -84.6623
1.7288 | -84.6568
1.7291 | -84.6521
1.7294 | -84.6449
1.7297 | -84.6393
1.73 | -84.6339
1.7303 | -84.6281
1.7306 | -84.6191
1.7309 | -84.6129
1.7312 | -84.6059
1.7316 | -84.5997
1.7319 | -84.591
1.7322 | -84.5821
1.7325 | -84.5756
1.7328 | -84.5665
1.7331 | -84.5608
1.7334 | -84.5556
1.7337 | -84.5491
1.734 | -84.5426
1.7344 | -84.5367
1.7347 | -84.5306
1.735 | -84.5259
1.7353 | -84.5199
1.7356 | -84.5139
1.7359 | -84.5084
1.7362 | -84.5028
1.7365 | -84.4964
1.7368 | -84.4908
1.7371 | -84.4872
1.7375 | -84.4863
1.7378 | -84.4815
1.7381 | -84.4796
1.7384 | -84.4742
1.7387 | -84.4701
1.739 | -84.4676
1.7393 | -84.4635
1.7396 | -84.462
1.7399 | -84.4637
1.7403 | -84.4624
1.7406 | -84.4592
1.7409 | -84.4578
1.7412 | -84.4544
1.7415 | -84.4461
1.7418 | -84.4407
1.7421 | -84.4318
1.7424 | -84.4218
1.7427 | -84.4164
1.743 | -84.4129
1.7434 | -84.4099
1.7437 | -84.4049
1.744 | -84.4007
1.7443 | -84.3959
1.7446 | -84.3923
1.7449 | -84.389
1.7452 | -84.3849
1.7455 | -84.3806
1.7458 | -84.3766
1.7462 | -84.3729
1.7465 | -84.3705
1.7468 | -84.3674
1.7471 | -84.3693
1.7474 | -84.3664
1.7477 | -84.364
1.748 | -84.361
1.7483 | -84.3563
1.7486 | -84.3552
1.7489 | -84.3541
1.7493 | -84.3516
1.7496 | -84.3503
1.7499 | -84.3489
1.7502 | -84.3458
1.7505 | -84.3395
1.7508 | -84.3342
1.7511 | -84.3321
1.7514 | -84.3297
1.7517 | -84.3274
1.7521 | -84.3254
1.7524 | -84.3216
1.7527 | -84.3205
1.753 | -84.3211
1.7533 | -84.3197
1.7536 | -84.3201
1.7539 | -84.3226
1.7542 | -84.3212
1.7545 | -84.3201
1.7548 | -84.3192
1.7552 | -84.3175
1.7555 | -84.3161
1.7558 | -84.3172
1.7561 | -84.3214
1.7564 | -84.3263
1.7567 | -84.3256
1.757 | -84.3252
1.7573 | -84.3248
1.7576 | -84.3261
1.758 | -84.3284
1.7583 | -84.329
1.7586 | -84.3297
1.7589 | -84.3306
1.7592 | -84.3312
1.7595 | -84.3323
1.7598 | -84.3314
1.7601 | -84.3314
1.7604 | -84.3339
1.7607 | -84.336
1.7611 | -84.3372
1.7614 | -84.338
1.7617 | -84.3386
1.762 | -84.3407
1.7623 | -84.3406
1.7626 | -84.342
1.7629 | -84.3427
1.7632 | -84.3449
1.7635 | -84.3464
1.7639 | -84.3486
1.7642 | -84.3503
1.7645 | -84.3502
1.7648 | -84.3537
1.7651 | -84.3565
1.7654 | -84.3595
1.7657 | -84.3642
1.766 | -84.3679
1.7663 | -84.3725
1.7666 | -84.3756
1.767 | -84.3788
1.7673 | -84.3809
1.7676 | -84.383
1.7679 | -84.3872
1.7682 | -84.3891
1.7685 | -84.3936
1.7688 | -84.3949
1.7691 | -84.3987
1.7694 | -84.4014
1.7698 | -84.3995
1.7701 | -84.4024
1.7704 | -84.4069
1.7707 | -84.4116
1.771 | -84.4147
1.7713 | -84.4179
1.7716 | -84.4206
1.7719 | -84.4224
1.7722 | -84.4273
1.7726 | -84.4335
1.7729 | -84.439
1.7732 | -84.4418
1.7735 | -84.4458
1.7738 | -84.4515
1.7741 | -84.4552
1.7744 | -84.4589
1.7747 | -84.4627
1.775 | -84.4666
1.7753 | -84.4707
1.7757 | -84.4738
1.776 | -84.4785
1.7763 | -84.481
1.7766 | -84.4867
1.7769 | -84.4916
1.7772 | -84.4971
1.7775 | -84.4991
1.7778 | -84.5033
1.7781 | -84.5102
1.7785 | -84.515
1.7788 | -84.5197
1.7791 | -84.5233
1.7794 | -84.5282
1.7797 | -84.5307
1.78 | -84.5301
1.7803 | -84.5294
1.7806 | -84.5318
1.7809 | -84.5343
1.7812 | -84.5363
1.7816 | -84.538
1.7819 | -84.5403
1.7822 | -84.5434
1.7825 | -84.5477
1.7828 | -84.5518
1.7831 | -84.5555
1.7834 | -84.5651
1.7837 | -84.5729
1.784 | -84.5747
1.7844 | -84.5738
1.7847 | -84.574
1.785 | -84.5745
1.7853 | -84.5734
1.7856 | -84.575
1.7859 | -84.579
1.7862 | -84.5831
1.7865 | -84.5863
1.7868 | -84.5891
1.7871 | -84.592
1.7875 | -84.5954
1.7878 | -84.5985
1.7881 | -84.6022
1.7884 | -84.6051
1.7887 | -84.6075
1.789 | -84.6108
1.7893 | -84.6137
1.7896 | -84.6152
1.7899 | -84.6176
1.7902 | -84.6227
1.7906 | -84.6246
1.7909 | -84.6237
1.7912 | -84.6257
1.7915 | -84.6289
1.7918 | -84.63
1.7921 | -84.6321
1.7924 | -84.6338
1.7927 | -84.6367
1.793 | -84.6378
1.7934 | -84.64
1.7937 | -84.6428
1.794 | -84.6425
1.7943 | -84.6441
1.7946 | -84.645
1.7949 | -84.6451
1.7952 | -84.6445
1.7955 | -84.6421
1.7958 | -84.6437
1.7962 | -84.6443
1.7965 | -84.647
1.7968 | -84.6457
1.7971 | -84.644
1.7974 | -84.6414
1.7977 | -84.6406
1.798 | -84.6389
1.7983 | -84.6376
1.7986 | -84.6371
1.7989 | -84.6359
1.7993 | -84.6374
1.7996 | -84.637
1.7999 | -84.638
1.8002 | -84.6384
1.8005 | -84.6401
1.8008 | -84.6384
1.8011 | -84.6388
1.8014 | -84.6372
1.8017 | -84.633
1.8021 | -84.63
1.8024 | -84.6287
1.8027 | -84.6292
1.803 | -84.6287
1.8033 | -84.6267
1.8036 | -84.6246
1.8039 | -84.6228
1.8042 | -84.6211
1.8045 | -84.6193
1.8048 | -84.6167
1.8052 | -84.6152
1.8055 | -84.6138
1.8058 | -84.6121
1.8061 | -84.6099
1.8064 | -84.6083
1.8067 | -84.6061
1.807 | -84.6014
1.8073 | -84.5981
1.8076 | -84.5957
1.808 | -84.5945
1.8083 | -84.5914
1.8086 | -84.5909
1.8089 | -84.5893
1.8092 | -84.5831
1.8095 | -84.5794
1.8098 | -84.5699
1.8101 | -84.5659
1.8104 | -84.5625
1.8107 | -84.5596
1.8111 | -84.557
1.8114 | -84.554
1.8117 | -84.5523
1.812 | -84.5518
1.8123 | -84.5502
1.8126 | -84.5448
1.8129 | -84.5392
1.8132 | -84.5346
1.8135 | -84.5294
1.8139 | -84.5233
1.8142 | -84.5214
1.8145 | -84.5201
1.8148 | -84.5147
1.8151 | -84.5119
1.8154 | -84.5072
1.8157 | -84.503
1.816 | -84.5011
1.8163 | -84.4994
1.8166 | -84.4947
1.817 | -84.4918
1.8173 | -84.4889
1.8176 | -84.4866
1.8179 | -84.4844
1.8182 | -84.4864
1.8185 | -84.4887
1.8188 | -84.4833
1.8191 | -84.4792
1.8194 | -84.4749
1.8198 | -84.4723
1.8201 | -84.4673
1.8204 | -84.4656
1.8207 | -84.4622
1.821 | -84.4572
1.8213 | -84.4527
1.8216 | -84.4456
1.8219 | -84.4405
1.8222 | -84.4354
1.8225 | -84.4296
1.8229 | -84.4285
1.8232 | -84.4289
1.8235 | -84.4264
1.8238 | -84.4187
1.8241 | -84.4113
1.8244 | -84.404
1.8247 | -84.3958
1.825 | -84.3854
1.8253 | -84.3776
1.8257 | -84.3712
1.826 | -84.3656
1.8263 | -84.3626
1.8266 | -84.3581
1.8269 | -84.3522
1.8272 | -84.3462
1.8275 | -84.3425
1.8278 | -84.337
1.8281 | -84.3309
1.8284 | -84.3223
1.8288 | -84.3143
1.8291 | -84.3079
1.8294 | -84.3026
1.8297 | -84.2973
1.83 | -84.2915
1.8303 | -84.2862
1.8306 | -84.2809
1.8309 | -84.2759
1.8312 | -84.271
1.8316 | -84.2662
1.8319 | -84.2608
1.8322 | -84.2557
1.8325 | -84.2506
1.8328 | -84.2453
1.8331 | -84.2393
1.8334 | -84.2329
1.8337 | -84.2257
1.834 | -84.2193
1.8343 | -84.2147
1.8347 | -84.2099
1.835 | -84.2052
1.8353 | -84.1999
1.8356 | -84.19
1.8359 | -84.1883
1.8362 | -84.1885
1.8365 | -84.1832
1.8368 | -84.1803
1.8371 | -84.1759
1.8375 | -84.1722
1.8378 | -84.1677
1.8381 | -84.1627
1.8384 | -84.1599
1.8387 | -84.1562
1.839 | -84.1529
1.8393 | -84.1487
1.8396 | -84.1457
1.8399 | -84.1421
1.8402 | -84.1384
1.8406 | -84.1355
1.8409 | -84.1314
1.8412 | -84.1282
1.8415 | -84.1251
1.8418 | -84.1222
1.8421 | -84.1212
1.8424 | -84.119
1.8427 | -84.1171
1.843 | -84.1148
1.8434 | -84.1122
1.8437 | -84.1102
1.844 | -84.1102
1.8443 | -84.1075
1.8446 | -84.1045
1.8449 | -84.103
1.8452 | -84.0993
1.8455 | -84.0998
1.8458 | -84.0965
1.8461 | -84.0941
1.8465 | -84.0938
1.8468 | -84.093
1.8471 | -84.0898
1.8474 | -84.0882
1.8477 | -84.0873
1.848 | -84.0871
1.8483 | -84.0866
1.8486 | -84.0843
1.8489 | -84.0813
1.8493 | -84.0808
1.8496 | -84.081
1.8499 | -84.0809
1.8502 | -84.0801
1.8505 | -84.0797
1.8508 | -84.0799
1.8511 | -84.0788
1.8514 | -84.0798
1.8517 | -84.0815
1.852 | -84.083
1.8524 | -84.0825
1.8527 | -84.0819
1.853 | -84.0791
1.8533 | -84.0761
1.8536 | -84.0758
1.8539 | -84.0747
1.8542 | -84.0724
1.8545 | -84.0747
1.8548 | -84.0754
1.8552 | -84.0763
1.8555 | -84.0766
1.8558 | -84.0786
1.8561 | -84.0774
1.8564 | -84.0799
1.8567 | -84.0818
1.857 | -84.0849
1.8573 | -84.0877
1.8576 | -84.0887
1.8579 | -84.0899
1.8583 | -84.089
1.8586 | -84.0837
1.8589 | -84.0834
1.8592 | -84.0861
1.8595 | -84.0879
1.8598 | -84.0895
1.8601 | -84.0895
1.8604 | -84.0855
1.8607 | -84.0831
1.8611 | -84.0841
1.8614 | -84.0852
1.8617 | -84.0863
1.862 | -84.0876
1.8623 | -84.0892
1.8626 | -84.09
1.8629 | -84.091
1.8632 | -84.0919
1.8635 | -84.0914
1.8638 | -84.0912
1.8642 | -84.0928
1.8645 | -84.0947
1.8648 | -84.0969
1.8651 | -84.1
1.8654 | -84.1021
1.8657 | -84.1037
1.866 | -84.1062
1.8663 | -84.1058
1.8666 | -84.0982
1.867 | -84.0967
1.8673 | -84.1003
1.8676 | -84.1025
1.8679 | -84.1076
1.8682 | -84.1149
1.8685 | -84.1191
1.8688 | -84.1208
1.8691 | -84.1222
1.8694 | -84.1245
1.8697 | -84.1252
1.8701 | -84.1262
1.8704 | -84.1256
1.8707 | -84.124
1.871 | -84.1245
1.8713 | -84.1276
1.8716 | -84.1284
1.8719 | -84.1311
1.8722 | -84.1343
1.8725 | -84.1323
1.8729 | -84.1315
1.8732 | -84.1336
1.8735 | -84.1347
1.8738 | -84.1378
1.8741 | -84.1393
1.8744 | -84.1408
1.8747 | -84.1454
1.875 | -84.1475
1.8753 | -84.1497
1.8756 | -84.151
1.876 | -84.1538
1.8763 | -84.1548
1.8766 | -84.1563
1.8769 | -84.1627
1.8772 | -84.1683
1.8775 | -84.1709
1.8778 | -84.1744
1.8781 | -84.1776
1.8784 | -84.1803
1.8788 | -84.1846
1.8791 | -84.186
1.8794 | -84.1883
1.8797 | -84.1908
1.88 | -84.1932
1.8803 | -84.1957
1.8806 | -84.1994
1.8809 | -84.2028
1.8812 | -84.2038
1.8815 | -84.1985
1.8819 | -84.2003
1.8822 | -84.2026
1.8825 | -84.2067
1.8828 | -84.2099
1.8831 | -84.2121
1.8834 | -84.2143
1.8837 | -84.2164
1.884 | -84.2184
1.8843 | -84.2211
1.8847 | -84.2232
1.885 | -84.2254
1.8853 | -84.2274
1.8856 | -84.2298
1.8859 | -84.2311
1.8862 | -84.2334
1.8865 | -84.2356
1.8868 | -84.2393
1.8871 | -84.2406
1.8874 | -84.2424
1.8878 | -84.2457
1.8881 | -84.2476
1.8884 | -84.2475
1.8887 | -84.2493
1.889 | -84.2509
1.8893 | -84.2519
1.8896 | -84.2496
1.8899 | -84.2507
1.8902 | -84.2499
1.8906 | -84.2505
1.8909 | -84.249
1.8912 | -84.2477
1.8915 | -84.2492
1.8918 | -84.2526
1.8921 | -84.2553
1.8924 | -84.2568
1.8927 | -84.2544
1.893 | -84.2533
1.8933 | -84.254
1.8937 | -84.2534
1.894 | -84.2514
1.8943 | -84.2511
1.8946 | -84.252
1.8949 | -84.253
1.8952 | -84.2533
1.8955 | -84.2516
1.8958 | -84.2522
1.8961 | -84.2508
1.8965 | -84.2506
1.8968 | -84.2498
1.8971 | -84.2516
1.8974 | -84.2569
1.8977 | -84.2572
1.898 | -84.251
1.8983 | -84.2427
1.8986 | -84.2416
1.8989 | -84.2394
1.8992 | -84.2253
1.8996 | -84.224
1.8999 | -84.2241
1.9002 | -84.2221
1.9005 | -84.2191
1.9008 | -84.2163
1.9011 | -84.215
1.9014 | -84.2114
1.9017 | -84.2075
1.902 | -84.2043
1.9024 | -84.2018
1.9027 | -84.1988
1.903 | -84.1957
1.9033 | -84.1931
1.9036 | -84.1905
1.9039 | -84.1895
1.9042 | -84.1871
1.9045 | -84.1825
1.9048 | -84.1704
1.9051 | -84.1632
1.9055 | -84.1587
1.9058 | -84.1574
1.9061 | -84.1533
1.9064 | -84.1453
1.9067 | -84.1394
1.907 | -84.1358
1.9073 | -84.1323
1.9076 | -84.1284
1.9079 | -84.1264
1.9083 | -84.1193
1.9086 | -84.1172
1.9089 | -84.114
1.9092 | -84.1091
1.9095 | -84.1057
1.9098 | -84.1017
1.9101 | -84.0972
1.9104 | -84.0929
1.9107 | -84.0883
1.911 | -84.0829
1.9114 | -84.0762
1.9117 | -84.0696
1.912 | -84.0644
1.9123 | -84.0586
1.9126 | -84.0535
1.9129 | -84.0501
1.9132 | -84.0447
1.9135 | -84.0423
1.9138 | -84.0381
1.9142 | -84.0349
1.9145 | -84.0293
1.9148 | -84.0223
1.9151 | -84.0159
1.9154 | -84.0104
1.9157 | -84.0044
1.916 | -83.9986
1.9163 | -83.992
1.9166 | -83.9842
1.9169 | -83.9778
1.9173 | -83.9701
1.9176 | -83.962
1.9179 | -83.956
1.9182 | -83.9495
1.9185 | -83.9432
1.9188 | -83.9363
1.9191 | -83.9297
1.9194 | -83.9237
1.9197 | -83.9181
1.9201 | -83.9147
1.9204 | -83.9084
1.9207 | -83.9024
1.921 | -83.896
1.9213 | -83.8874
1.9216 | -83.8786
1.9219 | -83.8723
1.9222 | -83.868
1.9225 | -83.8623
1.9228 | -83.8562
1.9232 | -83.8499
1.9235 | -83.8441
1.9238 | -83.8376
1.9241 | -83.833
1.9244 | -83.828
1.9247 | -83.8232
1.925 | -83.8186
1.9253 | -83.8112
1.9256 | -83.8057
1.926 | -83.7985
1.9263 | -83.7921
1.9266 | -83.7868
1.9269 | -83.7819
1.9272 | -83.7773
1.9275 | -83.771
1.9278 | -83.7636
1.9281 | -83.7568
1.9284 | -83.7506
1.9287 | -83.7461
1.9291 | -83.74
1.9294 | -83.7354
1.9297 | -83.7291
1.93 | -83.7224
1.9303 | -83.715
1.9306 | -83.7089
1.9309 | -83.702
1.9312 | -83.6935
1.9315 | -83.6861
1.9319 | -83.6826
1.9322 | -83.6799
1.9325 | -83.6741
1.9328 | -83.6669
1.9331 | -83.6596
1.9334 | -83.6541
1.9337 | -83.6492
1.934 | -83.6441
1.9343 | -83.6394
1.9346 | -83.6349
1.935 | -83.6302
1.9353 | -83.6273
1.9356 | -83.6214
1.9359 | -83.6152
1.9362 | -83.6079
1.9365 | -83.6027
1.9368 | -83.5966
1.9371 | -83.591
1.9374 | -83.5877
1.9378 | -83.5826
1.9381 | -83.5759
1.9384 | -83.5667
1.9387 | -83.5628
1.939 | -83.5546
1.9393 | -83.5472
1.9396 | -83.5471
1.9399 | -83.5425
1.9402 | -83.539
1.9406 | -83.5348
1.9409 | -83.5295
1.9412 | -83.5233
1.9415 | -83.5178
1.9418 | -83.5138
1.9421 | -83.5066
1.9424 | -83.5007
1.9427 | -83.4964
1.943 | -83.4919
1.9433 | -83.4882
1.9437 | -83.4828
1.944 | -83.4803
1.9443 | -83.4764
1.9446 | -83.4727
1.9449 | -83.47
1.9452 | -83.4658
1.9455 | -83.4601
1.9458 | -83.4557
1.9461 | -83.4503
1.9465 | -83.4468
1.9468 | -83.4433
1.9471 | -83.4372
1.9474 | -83.4331
1.9477 | -83.4297
1.948 | -83.4252
1.9483 | -83.4212
1.9486 | -83.4173
1.9489 | -83.4132
1.9492 | -83.4104
1.9496 | -83.4068
1.9499 | -83.4012
1.9502 | -83.396
1.9505 | -83.3908
1.9508 | -83.3869
1.9511 | -83.3832
1.9514 | -83.3794
1.9517 | -83.3758
1.952 | -83.3709
1.9524 | -83.3677
1.9527 | -83.3647
1.953 | -83.3651
1.9533 | -83.3637
1.9536 | -83.3602
1.9539 | -83.3577
1.9542 | -83.3563
1.9545 | -83.3547
1.9548 | -83.3531
1.9551 | -83.3519
1.9555 | -83.3518
1.9558 | -83.3521
1.9561 | -83.3519
1.9564 | -83.3518
1.9567 | -83.3535
1.957 | -83.3543
1.9573 | -83.3537
1.9576 | -83.3517
1.9579 | -83.35
1.9583 | -83.3487
1.9586 | -83.3463
1.9589 | -83.3441
1.9592 | -83.3413
1.9595 | -83.3396
1.9598 | -83.338
1.9601 | -83.3364
1.9604 | -83.3358
1.9607 | -83.3363
1.961 | -83.3357
1.9614 | -83.3337
1.9617 | -83.3307
1.962 | -83.3319
1.9623 | -83.3314
1.9626 | -83.3305
1.9629 | -83.3297
1.9632 | -83.3292
1.9635 | -83.3287
1.9638 | -83.327
1.9642 | -83.3258
1.9645 | -83.3261
1.9648 | -83.3261
1.9651 | -83.3261
1.9654 | -83.3278
1.9657 | -83.3298
1.966 | -83.332
1.9663 | -83.3336
1.9666 | -83.3349
1.9669 | -83.3353
1.9673 | -83.3361
1.9676 | -83.3372
1.9679 | -83.3421
1.9682 | -83.343
1.9685 | -83.3441
1.9688 | -83.345
1.9691 | -83.3442
1.9694 | -83.344
1.9697 | -83.3448
1.9701 | -83.3453
1.9704 | -83.3447
1.9707 | -83.3461
1.971 | -83.3506
1.9713 | -83.3519
1.9716 | -83.3523
1.9719 | -83.3541
1.9722 | -83.3555
1.9725 | -83.3553
1.9728 | -83.3563
1.9732 | -83.3582
1.9735 | -83.3599
1.9738 | -83.3619
1.9741 | -83.363
1.9744 | -83.3642
1.9747 | -83.3663
1.975 | -83.3683
1.9753 | -83.3707
1.9756 | -83.3719
1.976 | -83.3733
1.9763 | -83.3761
1.9766 | -83.3785
1.9769 | -83.3808
1.9772 | -83.3832
1.9775 | -83.386
1.9778 | -83.3859
1.9781 | -83.3853
1.9784 | -83.3876
1.9787 | -83.3902
1.9791 | -83.394
1.9794 | -83.396
1.9797 | -83.3994
1.98 | -83.3999
1.9803 | -83.4031
1.9806 | -83.4064
1.9809 | -83.4092
1.9812 | -83.4103
1.9815 | -83.4083
1.9819 | -83.4096
1.9822 | -83.4135
1.9825 | -83.4178
1.9828 | -83.4204
1.9831 | -83.4229
1.9834 | -83.4258
1.9837 | -83.4294
1.984 | -83.4322
1.9843 | -83.4358
1.9846 | -83.44
1.985 | -83.4445
1.9853 | -83.448
1.9856 | -83.4516
1.9859 | -83.4549
1.9862 | -83.459
1.9865 | -83.4649
1.9868 | -83.4708
1.9871 | -83.4744
1.9874 | -83.4792
1.9878 | -83.4859
1.9881 | -83.4914
1.9884 | -83.4937
1.9887 | -83.4977
1.989 | -83.5038
1.9893 | -83.508
1.9896 | -83.5122
1.9899 | -83.5172
1.9902 | -83.521
1.9905 | -83.5241
1.9909 | -83.5279
1.9912 | -83.5306
1.9915 | -83.5341
1.9918 | -83.5377
1.9921 | -83.5438
1.9924 | -83.5503
1.9927 | -83.5597
1.993 | -83.5659
1.9933 | -83.5681
1.9937 | -83.5701
1.994 | -83.5738
1.9943 | -83.578
1.9946 | -83.5831
1.9949 | -83.585
1.9952 | -83.5887
1.9955 | -83.5929
1.9958 | -83.5975
1.9961 | -83.6013
1.9964 | -83.6052
1.9968 | -83.61
1.9971 | -83.614
1.9974 | -83.6189
1.9977 | -83.6255
1.998 | -83.6308
1.9983 | -83.6378
1.9986 | -83.6439
1.9989 | -83.648
1.9992 | -83.6524
1.9996 | -83.6566
1.9999 | -83.6593
2.0002 | -83.6631
2.0005 | -83.6649
2.0008 | -83.6674
2.0011 | -83.6742
2.0014 | -83.6865
2.0017 | -83.6919
2.002 | -83.6931
2.0023 | -83.697
2.0027 | -83.6999
2.003 | -83.7001
2.0033 | -83.7027
2.0036 | -83.7078
2.0039 | -83.7123
2.0042 | -83.7156
2.0045 | -83.7206
2.0048 | -83.7257
2.0051 | -83.7314
2.0055 | -83.7379
2.0058 | -83.7415
2.0061 | -83.7453
2.0064 | -83.7506
2.0067 | -83.755
2.007 | -83.7566
2.0073 | -83.76
2.0076 | -83.7631
2.0079 | -83.7717
2.0082 | -83.7758
2.0086 | -83.7773
2.0089 | -83.7819
2.0092 | -83.7836
2.0095 | -83.7883
2.0098 | -83.7936
2.0101 | -83.7967
2.0104 | -83.7991
2.0107 | -83.8016
2.011 | -83.8061
2.0114 | -83.8081
2.0117 | -83.8104
2.012 | -83.8138
2.0123 | -83.8166
2.0126 | -83.8153
2.0129 | -83.8177
2.0132 | -83.8203
2.0135 | -83.824
2.0138 | -83.8267
2.0141 | -83.8305
2.0145 | -83.8352
2.0148 | -83.838
2.0151 | -83.8411
2.0154 | -83.8442
2.0157 | -83.8457
2.016 | -83.8487
2.0163 | -83.8518
2.0166 | -83.8556
2.0169 | -83.8589
2.0173 | -83.8631
2.0176 | -83.869
2.0179 | -83.8734
2.0182 | -83.8768
2.0185 | -83.8799
2.0188 | -83.8826
2.0191 | -83.8836
2.0194 | -83.8805
2.0197 | -83.8811
2.02 | -83.8846
2.0204 | -83.8852
2.0207 | -83.8873
2.021 | -83.8876
2.0213 | -83.8901
2.0216 | -83.8931
2.0219 | -83.8962
2.0222 | -83.8981
2.0225 | -83.8997
2.0228 | -83.9021
2.0232 | -83.9057
2.0235 | -83.9087
2.0238 | -83.9092
2.0241 | -83.9095
2.0244 | -83.9137
2.0247 | -83.9133
2.025 | -83.917
2.0253 | -83.9197
2.0256 | -83.9207
2.0259 | -83.9248
2.0263 | -83.9266
2.0266 | -83.9291
2.0269 | -83.9308
2.0272 | -83.9325
2.0275 | -83.9369
2.0278 | -83.9407
2.0281 | -83.9406
2.0284 | -83.9391
2.0287 | -83.9407
2.0291 | -83.9416
2.0294 | -83.9442
2.0297 | -83.9453
2.03 | -83.9451
2.0303 | -83.9465
2.0306 | -83.9488
2.0309 | -83.9505
2.0312 | -83.9512
2.0315 | -83.9512
2.0318 | -83.9523
2.0322 | -83.954
2.0325 | -83.9547
2.0328 | -83.9565
2.0331 | -83.9601
2.0334 | -83.9627
2.0337 | -83.9631
2.034 | -83.9627
2.0343 | -83.9634
2.0346 | -83.9642
2.035 | -83.9693
2.0353 | -83.9752
2.0356 | -83.9749
2.0359 | -83.977
2.0362 | -83.9825
2.0365 | -83.9875
2.0368 | -83.9881
2.0371 | -83.9874
2.0374 | -83.9876
2.0377 | -83.986
2.0381 | -83.9866
2.0384 | -83.9862
2.0387 | -83.9874
2.039 | -83.9894
2.0393 | -83.9893
2.0396 | -83.9893
2.0399 | -83.99
2.0402 | -83.9897
2.0405 | -83.988
2.0409 | -83.9872
2.0412 | -83.9854
2.0415 | -83.9849
2.0418 | -83.9869
2.0421 | -83.9912
2.0424 | -83.9942
2.0427 | -83.9937
2.043 | -83.9921
2.0433 | -83.9908
2.0436 | -83.9904
2.044 | -83.9894
2.0443 | -83.9885
2.0446 | -83.9879
2.0449 | -83.9867
2.0452 | -83.9862
2.0455 | -83.9849
2.0458 | -83.9832
2.0461 | -83.9828
2.0464 | -83.9824
2.0468 | -83.9839
2.0471 | -83.9845
2.0474 | -83.9815
2.0477 | -83.9792
2.048 | -83.9778
2.0483 | -83.9747
2.0486 | -83.9686
2.0489 | -83.9661
2.0492 | -83.9636
2.0495 | -83.9638
2.0499 | -83.9613
2.0502 | -83.9596
2.0505 | -83.9581
2.0508 | -83.9556
2.0511 | -83.9572
2.0514 | -83.9534
2.0517 | -83.9518
2.052 | -83.9509
2.0523 | -83.9496
2.0527 | -83.9474
2.053 | -83.9464
2.0533 | -83.9433
2.0536 | -83.9393
2.0539 | -83.9378
2.0542 | -83.9343
2.0545 | -83.9319
2.0548 | -83.9314
2.0551 | -83.9298
2.0554 | -83.9276
2.0558 | -83.9261
2.0561 | -83.9221
2.0564 | -83.9155
2.0567 | -83.914
2.057 | -83.9107
2.0573 | -83.9085
2.0576 | -83.9061
2.0579 | -83.9035
2.0582 | -83.9017
2.0586 | -83.8995
2.0589 | -83.8961
2.0592 | -83.8914
2.0595 | -83.8875
2.0598 | -83.8846
2.0601 | -83.8826
2.0604 | -83.8789
2.0607 | -83.871
2.061 | -83.8678
2.0614 | -83.8647
2.0617 | -83.8614
2.062 | -83.8516
2.0623 | -83.8494
2.0626 | -83.846
2.0629 | -83.8438
2.0632 | -83.8382
2.0635 | -83.8356
2.0638 | -83.8317
2.0641 | -83.828
2.0645 | -83.8231
2.0648 | -83.819
2.0651 | -83.8183
2.0654 | -83.8152
2.0657 | -83.8127
2.066 | -83.8104
2.0663 | -83.8073
2.0666 | -83.8057
2.0669 | -83.8047
2.0673 | -83.802
2.0676 | -83.7995
2.0679 | -83.7966
2.0682 | -83.7938
2.0685 | -83.7904
2.0688 | -83.7859
2.0691 | -83.7824
2.0694 | -83.7801
2.0697 | -83.7772
2.07 | -83.7727
2.0704 | -83.7664
2.0707 | -83.7633
2.071 | -83.7625
2.0713 | -83.7594
2.0716 | -83.7565
2.0719 | -83.7531
2.0722 | -83.7509
2.0725 | -83.7473
2.0728 | -83.7393
2.0732 | -83.7352
2.0735 | -83.7316
2.0738 | -83.7273
2.0741 | -83.7229
2.0744 | -83.7173
2.0747 | -83.7108
2.075 | -83.7079
2.0753 | -83.7053
2.0756 | -83.7035
2.0759 | -83.7025
2.0763 | -83.7006
2.0766 | -83.6988
2.0769 | -83.696
2.0772 | -83.6928
2.0775 | -83.6908
2.0778 | -83.6894
2.0781 | -83.6911
2.0784 | -83.6881
2.0787 | -83.6848
2.0791 | -83.6852
2.0794 | -83.6849
2.0797 | -83.6826
2.08 | -83.6796
2.0803 | -83.6768
2.0806 | -83.6746
2.0809 | -83.6722
2.0812 | -83.6695
2.0815 | -83.6645
2.0818 | -83.6581
2.0822 | -83.6567
2.0825 | -83.6554
2.0828 | -83.6514
2.0831 | -83.6472
2.0834 | -83.6435
2.0837 | -83.6416
2.084 | -83.6392
2.0843 | -83.6332
2.0846 | -83.6266
2.085 | -83.6221
2.0853 | -83.6177
2.0856 | -83.6134
2.0859 | -83.6119
2.0862 | -83.6103
2.0865 | -83.6067
2.0868 | -83.6002
2.0871 | -83.5959
2.0874 | -83.5933
2.0877 | -83.5909
2.0881 | -83.5859
2.0884 | -83.5837
2.0887 | -83.5797
2.089 | -83.5737
2.0893 | -83.5696
2.0896 | -83.5658
2.0899 | -83.5619
2.0902 | -83.5581
2.0905 | -83.5557
2.0909 | -83.5518
2.0912 | -83.5486
2.0915 | -83.5447
2.0918 | -83.538
2.0921 | -83.5313
2.0924 | -83.5285
2.0927 | -83.5262
2.093 | -83.5225
2.0933 | -83.5183
2.0936 | -83.5139
2.094 | -83.51
2.0943 | -83.5051
2.0946 | -83.5028
2.0949 | -83.4996
2.0952 | -83.4944
2.0955 | -83.4876
2.0958 | -83.4816
2.0961 | -83.4775
2.0964 | -83.4735
2.0968 | -83.4692
2.0971 | -83.4652
2.0974 | -83.4615
2.0977 | -83.4581
2.098 | -83.452
2.0983 | -83.4466
2.0986 | -83.4433
2.0989 | -83.4403
2.0992 | -83.4372
2.0995 | -83.4334
2.0999 | -83.4268
2.1002 | -83.4194
2.1005 | -83.4155
2.1008 | -83.4124
2.1011 | -83.4092
2.1014 | -83.4063
2.1017 | -83.4014
2.102 | -83.3939
2.1023 | -83.3886
2.1027 | -83.3847
2.103 | -83.3782
2.1033 | -83.3725
2.1036 | -83.3666
2.1039 | -83.3606
2.1042 | -83.3553
2.1045 | -83.3508
2.1048 | -83.3466
2.1051 | -83.3416
2.1054 | -83.3378
2.1058 | -83.3352
2.1061 | -83.3321
2.1064 | -83.3267
2.1067 | -83.3225
2.107 | -83.3191
2.1073 | -83.3148
2.1076 | -83.3103
2.1079 | -83.3063
2.1082 | -83.3025
2.1086 | -83.2961
2.1089 | -83.2915
2.1092 | -83.2859
2.1095 | -83.2811
2.1098 | -83.2764
2.1101 | -83.2726
2.1104 | -83.2685
2.1107 | -83.2653
2.111 | -83.2641
2.1113 | -83.2604
2.1117 | -83.2566
2.112 | -83.2525
2.1123 | -83.2492
2.1126 | -83.2455
2.1129 | -83.2421
2.1132 | -83.2382
2.1135 | -83.2337
2.1138 | -83.2311
2.1141 | -83.228
2.1145 | -83.2248
2.1148 | -83.2205
2.1151 | -83.2173
2.1154 | -83.2138
2.1157 | -83.2101
2.116 | -83.2069
2.1163 | -83.2031
2.1166 | -83.1997
2.1169 | -83.1973
2.1172 | -83.1938
2.1176 | -83.1908
2.1179 | -83.1876
2.1182 | -83.1815
2.1185 | -83.1743
2.1188 | -83.1699
2.1191 | -83.1666
2.1194 | -83.1635
2.1197 | -83.1609
2.12 | -83.1574
2.1204 | -83.1527
2.1207 | -83.1518
2.121 | -83.1496
2.1213 | -83.1468
2.1216 | -83.1434
2.1219 | -83.1403
2.1222 | -83.1362
2.1225 | -83.1335
2.1228 | -83.1275
2.1231 | -83.1206
2.1235 | -83.1153
2.1238 | -83.1116
2.1241 | -83.1059
2.1244 | -83.103
2.1247 | -83.1039
2.125 | -83.1012
2.1253 | -83.0952
2.1256 | -83.0911
2.1259 | -83.0881
2.1263 | -83.0848
2.1266 | -83.0819
2.1269 | -83.0771
2.1272 | -83.0735
2.1275 | -83.0697
2.1278 | -83.0667
2.1281 | -83.0637
2.1284 | -83.0602
2.1287 | -83.0561
2.129 | -83.0522
2.1294 | -83.047
2.1297 | -83.0439
2.13 | -83.0413
2.1303 | -83.0372
2.1306 | -83.0318
2.1309 | -83.0289
2.1312 | -83.0266
2.1315 | -83.0234
2.1318 | -83.0196
2.1322 | -83.0184
2.1325 | -83.0157
2.1328 | -83.0111
2.1331 | -83.0038
2.1334 | -82.9963
2.1337 | -82.9924
2.134 | -82.9914
2.1343 | -82.9879
2.1346 | -82.9847
2.1349 | -82.9834
2.1353 | -82.9819
2.1356 | -82.9806
2.1359 | -82.9784
2.1362 | -82.9752
2.1365 | -82.971
2.1368 | -82.9677
2.1371 | -82.965
2.1374 | -82.9651
2.1377 | -82.9633
2.1381 | -82.9618
2.1384 | -82.9615
2.1387 | -82.9632
2.139 | -82.9613
2.1393 | -82.9587
2.1396 | -82.9565
2.1399 | -82.9547
2.1402 | -82.9529
2.1405 | -82.9507
2.1408 | -82.9474
2.1412 | -82.9461
2.1415 | -82.9435
2.1418 | -82.9409
2.1421 | -82.9387
2.1424 | -82.9372
2.1427 | -82.9347
2.143 | -82.9321
2.1433 | -82.9295
2.1436 | -82.9272
2.144 | -82.926
2.1443 | -82.9243
2.1446 | -82.9227
2.1449 | -82.9233
2.1452 | -82.9219
2.1455 | -82.9193
2.1458 | -82.9175
2.1461 | -82.917
2.1464 | -82.9187
2.1467 | -82.9167
2.1471 | -82.9138
2.1474 | -82.9147
2.1477 | -82.9133
2.148 | -82.9113
2.1483 | -82.9095
2.1486 | -82.9084
2.1489 | -82.9073
2.1492 | -82.9066
2.1495 | -82.9062
2.1499 | -82.9055
2.1502 | -82.9059
2.1505 | -82.9064
2.1508 | -82.9031
2.1511 | -82.8996
2.1514 | -82.8966
2.1517 | -82.8952
2.152 | -82.896
2.1523 | -82.8945
2.1526 | -82.8928
2.153 | -82.8923
2.1533 | -82.8932
2.1536 | -82.8941
2.1539 | -82.8932
2.1542 | -82.8933
2.1545 | -82.8925
2.1548 | -82.891
2.1551 | -82.8885
2.1554 | -82.8859
2.1558 | -82.8843
2.1561 | -82.8828
2.1564 | -82.883
2.1567 | -82.8844
2.157 | -82.8848
2.1573 | -82.8836
2.1576 | -82.8857
2.1579 | -82.8855
2.1582 | -82.8836
2.1585 | -82.8817
2.1589 | -82.8795
2.1592 | -82.8797
2.1595 | -82.8789
2.1598 | -82.8783
2.1601 | -82.8779
2.1604 | -82.877
2.1607 | -82.8764
2.161 | -82.8767
2.1613 | -82.8767
2.1617 | -82.8764
2.162 | -82.876
2.1623 | -82.8747
2.1626 | -82.8755
2.1629 | -82.877
2.1632 | -82.8766
2.1635 | -82.8755
2.1638 | -82.8743
2.1641 | -82.8739
2.1644 | -82.8729
2.1648 | -82.8735
2.1651 | -82.8773
2.1654 | -82.8786
2.1657 | -82.8786
2.166 | -82.88
2.1663 | -82.8799
2.1666 | -82.8816
2.1669 | -82.8816
2.1672 | -82.8808
2.1676 | -82.88
2.1679 | -82.88
2.1682 | -82.8819
2.1685 | -82.8794
2.1688 | -82.879
2.1691 | -82.8798
2.1694 | -82.8792
2.1697 | -82.8799
2.17 | -82.8799
2.1703 | -82.8789
2.1707 | -82.8791
2.171 | -82.8813
2.1713 | -82.8815
2.1716 | -82.8821
2.1719 | -82.8841
2.1722 | -82.8856
2.1725 | -82.8857
2.1728 | -82.8834
2.1731 | -82.8835
2.1735 | -82.884
2.1738 | -82.8834
2.1741 | -82.8835
2.1744 | -82.8852
2.1747 | -82.8878
2.175 | -82.8883
2.1753 | -82.8892
2.1756 | -82.891
2.1759 | -82.8922
2.1762 | -82.8927
2.1766 | -82.8941
2.1769 | -82.8954
2.1772 | -82.899
2.1775 | -82.8987
2.1778 | -82.8982
2.1781 | -82.9
2.1784 | -82.9013
2.1787 | -82.9016
2.179 | -82.8987
2.1794 | -82.8971
2.1797 | -82.8978
2.18 | -82.8974
2.1803 | -82.8976
2.1806 | -82.8982
2.1809 | -82.9003
2.1812 | -82.9011
2.1815 | -82.9028
2.1818 | -82.9032
2.1821 | -82.9041
2.1825 | -82.9057
2.1828 | -82.9038
2.1831 | -82.9036
2.1834 | -82.9042
2.1837 | -82.9043
2.184 | -82.9061
2.1843 | -82.9069
2.1846 | -82.9065
2.1849 | -82.9056
2.1853 | -82.9065
2.1856 | -82.9072
2.1859 | -82.909
2.1862 | -82.9101
2.1865 | -82.9101
2.1868 | -82.9103
2.1871 | -82.9105
2.1874 | -82.9094
2.1877 | -82.9115
2.188 | -82.9135
2.1884 | -82.9152
2.1887 | -82.9164
2.189 | -82.9194
2.1893 | -82.9223
2.1896 | -82.9239
2.1899 | -82.925
2.1902 | -82.9252
2.1905 | -82.9281
2.1908 | -82.9282
2.1912 | -82.9296
2.1915 | -82.9337
2.1918 | -82.9363
2.1921 | -82.938
2.1924 | -82.9393
2.1927 | -82.9405
2.193 | -82.9405
2.1933 | -82.9424
2.1936 | -82.9446
2.1939 | -82.9464
2.1943 | -82.9468
2.1946 | -82.9486
2.1949 | -82.9425
2.1952 | -82.9436
2.1955 | -82.9472
2.1958 | -82.9493
2.1961 | -82.952
2.1964 | -82.9573
2.1967 | -82.9597
2.1971 | -82.9602
2.1974 | -82.9585
2.1977 | -82.9614
2.198 | -82.9623
2.1983 | -82.9629
2.1986 | -82.9648
2.1989 | -82.9653
2.1992 | -82.9675
2.1995 | -82.9693
2.1999 | -82.9712
2.2002 | -82.9732
2.2005 | -82.975
2.2008 | -82.9748
2.2011 | -82.975
2.2014 | -82.9779
2.2017 | -82.9808
2.202 | -82.9823
2.2023 | -82.9849
2.2026 | -82.9883
2.203 | -82.9895
2.2033 | -82.9896
2.2036 | -82.9917
2.2039 | -82.9917
2.2042 | -82.9943
2.2045 | -82.999
2.2048 | -83.0038
2.2051 | -83.0053
2.2054 | -83.0074
2.2058 | -83.0104
2.2061 | -83.0124
2.2064 | -83.0132
2.2067 | -83.0144
2.207 | -83.0156
2.2073 | -83.0168
2.2076 | -83.02
2.2079 | -83.0233
2.2082 | -83.0264
2.2085 | -83.0278
2.2089 | -83.0303
2.2092 | -83.0359
2.2095 | -83.0381
2.2098 | -83.0401
2.2101 | -83.0427
2.2104 | -83.0452
2.2107 | -83.0524
2.211 | -83.0531
2.2113 | -83.0539
2.2117 | -83.0554
2.212 | -83.0595
2.2123 | -83.0604
2.2126 | -83.0625
2.2129 | -83.0636
2.2132 | -83.0667
2.2135 | -83.0695
2.2138 | -83.0726
2.2141 | -83.073
2.2144 | -83.0724
2.2148 | -83.0729
2.2151 | -83.0745
2.2154 | -83.0756
2.2157 | -83.0761
2.216 | -83.0767
2.2163 | -83.0779
2.2166 | -83.0786
2.2169 | -83.0799
2.2172 | -83.0798
2.2176 | -83.0781
2.2179 | -83.0807
2.2182 | -83.0824
2.2185 | -83.0818
2.2188 | -83.0819
2.2191 | -83.0841
2.2194 | -83.0866
2.2197 | -83.0914
2.22 | -83.0942
2.2203 | -83.0938
2.2207 | -83.0954
2.221 | -83.0963
2.2213 | -83.0979
2.2216 | -83.0989
2.2219 | -83.0982
2.2222 | -83.0991
2.2225 | -83.1039
2.2228 | -83.106
2.2231 | -83.1073
2.2235 | -83.1083
2.2238 | -83.1098
2.2241 | -83.1126
2.2244 | -83.1143
2.2247 | -83.1151
2.225 | -83.1159
2.2253 | -83.1151
2.2256 | -83.1138
2.2259 | -83.1141
2.2262 | -83.1143
2.2266 | -83.1149
2.2269 | -83.1154
2.2272 | -83.1161
2.2275 | -83.1172
2.2278 | -83.1178
2.2281 | -83.1195
2.2284 | -83.1204
2.2287 | -83.1218
2.229 | -83.1233
2.2294 | -83.1244
2.2297 | -83.1254
2.23 | -83.1233
2.2303 | -83.1221
2.2306 | -83.1229
2.2309 | -83.1239
2.2312 | -83.1248
2.2315 | -83.125
2.2318 | -83.1254
2.2321 | -83.1285
2.2325 | -83.129
2.2328 | -83.1285
2.2331 | -83.1294
2.2334 | -83.1308
2.2337 | -83.1315
2.234 | -83.1325
2.2343 | -83.1335
2.2346 | -83.1366
2.2349 | -83.14
2.2353 | -83.1413
2.2356 | -83.1413
2.2359 | -83.1405
2.2362 | -83.141
2.2365 | -83.1413
2.2368 | -83.1433
2.2371 | -83.1461
2.2374 | -83.1476
2.2377 | -83.1482
2.238 | -83.1491
2.2384 | -83.1507
2.2387 | -83.1508
2.239 | -83.15
2.2393 | -83.1512
2.2396 | -83.1548
2.2399 | -83.1491
2.2402 | -83.1466
2.2405 | -83.1466
2.2408 | -83.1473
2.2412 | -83.1478
2.2415 | -83.1484
2.2418 | -83.1503
2.2421 | -83.1518
2.2424 | -83.1519
2.2427 | -83.1521
2.243 | -83.1551
2.2433 | -83.1552
2.2436 | -83.1559
2.2439 | -83.156
2.2443 | -83.156
2.2446 | -83.1544
2.2449 | -83.1513
2.2452 | -83.1476
2.2455 | -83.1459
2.2458 | -83.1456
2.2461 | -83.1454
2.2464 | -83.1463
2.2467 | -83.1469
2.2471 | -83.1468
2.2474 | -83.147
2.2477 | -83.1477
2.248 | -83.1482
2.2483 | -83.1479
2.2486 | -83.1469
2.2489 | -83.1453
2.2492 | -83.1438
2.2495 | -83.1448
2.2498 | -83.146
2.2502 | -83.1449
2.2505 | -83.1425
2.2508 | -83.1411
2.2511 | -83.1414
2.2514 | -83.1444
2.2517 | -83.1472
2.252 | -83.1476
2.2523 | -83.1467
2.2526 | -83.1454
2.253 | -83.1455
2.2533 | -83.1473
2.2536 | -83.1484
2.2539 | -83.148
2.2542 | -83.1486
2.2545 | -83.1516
2.2548 | -83.1543
2.2551 | -83.1551
2.2554 | -83.1551
2.2557 | -83.1557
2.2561 | -83.1534
2.2564 | -83.1526
2.2567 | -83.1527
2.257 | -83.1524
2.2573 | -83.1508
2.2576 | -83.1447
2.2579 | -83.1342
2.2582 | -83.1317
2.2585 | -83.1322
2.2589 | -83.1286
2.2592 | -83.1215
2.2595 | -83.1182
2.2598 | -83.118
2.2601 | -83.1169
2.2604 | -83.116
2.2607 | -83.117
2.261 | -83.1165
2.2613 | -83.1169
2.2616 | -83.1173
2.262 | -83.1174
2.2623 | -83.1162
2.2626 | -83.1145
2.2629 | -83.1145
2.2632 | -83.1145
2.2635 | -83.1148
2.2638 | -83.1142
2.2641 | -83.1127
2.2644 | -83.112
2.2648 | -83.1121
2.2651 | -83.1116
2.2654 | -83.1114
2.2657 | -83.1107
2.266 | -83.1096
2.2663 | -83.1085
2.2666 | -83.1086
2.2669 | -83.1084
2.2672 | -83.1088
2.2675 | -83.1098
2.2679 | -83.1103
2.2682 | -83.1097
2.2685 | -83.1097
2.2688 | -83.1092
2.2691 | -83.1094
2.2694 | -83.1098
2.2697 | -83.1079
2.27 | -83.107
2.2703 | -83.1045
2.2707 | -83.1031
2.271 | -83.1064
2.2713 | -83.1092
2.2716 | -83.1101
2.2719 | -83.1099
2.2722 | -83.1095
2.2725 | -83.1079
2.2728 | -83.107
2.2731 | -83.1076
2.2734 | -83.1084
2.2738 | -83.1078
2.2741 | -83.1067
2.2744 | -83.1063
2.2747 | -83.107
2.275 | -83.1068
2.2753 | -83.1046
2.2756 | -83.1043
2.2759 | -83.1092
2.2762 | -83.1122
2.2766 | -83.1134
2.2769 | -83.1136
2.2772 | -83.1119
2.2775 | -83.1132
2.2778 | -83.1127
2.2781 | -83.1116
2.2784 | -83.1105
2.2787 | -83.1108
2.279 | -83.1103
2.2793 | -83.109
2.2797 | -83.107
2.28 | -83.1035
2.2803 | -83.103
2.2806 | -83.103
2.2809 | -83.1021
2.2812 | -83.0993
2.2815 | -83.0971
2.2818 | -83.0961
2.2821 | -83.0957
2.2825 | -83.0942
2.2828 | -83.0937
2.2831 | -83.0927
2.2834 | -83.0915
2.2837 | -83.0919
2.284 | -83.0931
2.2843 | -83.0921
2.2846 | -83.0928
2.2849 | -83.0916
2.2852 | -83.0924
2.2856 | -83.0933
2.2859 | -83.0924
2.2862 | -83.0916
2.2865 | -83.0892
2.2868 | -83.0887
2.2871 | -83.0915
2.2874 | -83.0947
2.2877 | -83.0931
2.288 | -83.0905
2.2884 | -83.0899
2.2887 | -83.0886
2.289 | -83.0872
2.2893 | -83.0856
2.2896 | -83.0847
2.2899 | -83.0823
2.2902 | -83.0788
2.2905 | -83.076
2.2908 | -83.0749
2.2911 | -83.0758
2.2915 | -83.0749
2.2918 | -83.0716
2.2921 | -83.0683
2.2924 | -83.0673
2.2927 | -83.0664
2.293 | -83.0647
2.2933 | -83.0633
2.2936 | -83.0618
2.2939 | -83.0609
2.2943 | -83.0589
2.2946 | -83.0572
2.2949 | -83.0566
2.2952 | -83.0571
2.2955 | -83.0568
2.2958 | -83.056
2.2961 | -83.0539
2.2964 | -83.0502
2.2967 | -83.0491
2.297 | -83.0477
2.2974 | -83.0457
2.2977 | -83.0444
2.298 | -83.0417
2.2983 | -83.0364
2.2986 | -83.0333
2.2989 | -83.0331
2.2992 | -83.0323
2.2995 | -83.0307
2.2998 | -83.0283
2.3002 | -83.0277
2.3005 | -83.0263
2.3008 | -83.0247
2.3011 | -83.0221
2.3014 | -83.02
2.3017 | -83.014
2.302 | -83.0081
2.3023 | -83.0046
2.3026 | -82.9989
2.3029 | -82.9963
2.3033 | -82.9944
2.3036 | -82.9922
2.3039 | -82.9896
2.3042 | -82.9837
2.3045 | -82.9792
2.3048 | -82.9756
2.3051 | -82.9725
2.3054 | -82.9704
2.3057 | -82.9662
2.3061 | -82.9651
2.3064 | -82.9655
2.3067 | -82.9646
2.307 | -82.9636
2.3073 | -82.9607
2.3076 | -82.9582
2.3079 | -82.9569
2.3082 | -82.9544
2.3085 | -82.9521
2.3088 | -82.9494
2.3092 | -82.9471
2.3095 | -82.945
2.3098 | -82.9417
2.3101 | -82.9379
2.3104 | -82.9342
2.3107 | -82.9341
2.311 | -82.9329
2.3113 | -82.9297
2.3116 | -82.9257
2.312 | -82.923
2.3123 | -82.9202
2.3126 | -82.9175
2.3129 | -82.9157
2.3132 | -82.9133
2.3135 | -82.91
2.3138 | -82.9086
2.3141 | -82.9069
2.3144 | -82.9061
2.3147 | -82.9032
2.3151 | -82.9004
2.3154 | -82.8979
2.3157 | -82.8946
2.316 | -82.8917
2.3163 | -82.8896
2.3166 | -82.8863
2.3169 | -82.8839
2.3172 | -82.8816
2.3175 | -82.879
2.3179 | -82.8761
2.3182 | -82.8728
2.3185 | -82.8698
2.3188 | -82.8677
2.3191 | -82.866
2.3194 | -82.8645
2.3197 | -82.863
2.32 | -82.8603
2.3203 | -82.8568
2.3206 | -82.8551
2.321 | -82.8531
2.3213 | -82.851
2.3216 | -82.8479
2.3219 | -82.8446
2.3222 | -82.8415
2.3225 | -82.8397
2.3228 | -82.8367
2.3231 | -82.834
2.3234 | -82.8323
2.3238 | -82.8301
2.3241 | -82.8272
2.3244 | -82.8248
2.3247 | -82.8229
2.325 | -82.8224
2.3253 | -82.8205
2.3256 | -82.8172
2.3259 | -82.8145
2.3262 | -82.8134
2.3265 | -82.8129
2.3269 | -82.8123
2.3272 | -82.8091
2.3275 | -82.8063
2.3278 | -82.8048
2.3281 | -82.8028
2.3284 | -82.8
2.3287 | -82.7979
2.329 | -82.7962
2.3293 | -82.7949
2.3297 | -82.7928
2.33 | -82.7904
2.3303 | -82.7874
2.3306 | -82.7856
2.3309 | -82.7835
2.3312 | -82.7815
2.3315 | -82.7793
2.3318 | -82.7769
2.3321 | -82.7751
2.3324 | -82.7733
2.3328 | -82.7715
2.3331 | -82.769
2.3334 | -82.7662
2.3337 | -82.7643
2.334 | -82.7622
2.3343 | -82.7602
2.3346 | -82.7591
2.3349 | -82.7588
2.3352 | -82.7567
2.3356 | -82.7529
2.3359 | -82.7516
2.3362 | -82.7541
2.3365 | -82.7545
2.3368 | -82.7526
2.3371 | -82.752
2.3374 | -82.751
2.3377 | -82.749
2.338 | -82.7447
2.3384 | -82.7435
2.3387 | -82.7415
2.339 | -82.7399
2.3393 | -82.7363
2.3396 | -82.7346
2.3399 | -82.7327
2.3402 | -82.7328
2.3405 | -82.7312
2.3408 | -82.7302
2.3411 | -82.7297
2.3415 | -82.7281
2.3418 | -82.7271
2.3421 | -82.7289
2.3424 | -82.7295
2.3427 | -82.7283
2.343 | -82.7261
2.3433 | -82.7243
2.3436 | -82.7242
2.3439 | -82.7219
2.3443 | -82.7215
2.3446 | -82.7204
2.3449 | -82.7184
2.3452 | -82.7182
2.3455 | -82.7171
2.3458 | -82.7158
2.3461 | -82.7142
2.3464 | -82.7126
2.3467 | -82.7106
2.347 | -82.7087
2.3474 | -82.7094
2.3477 | -82.7077
2.348 | -82.7053
2.3483 | -82.7048
2.3486 | -82.7036
2.3489 | -82.7036
2.3492 | -82.7025
2.3495 | -82.7011
2.3498 | -82.7004
2.3502 | -82.6984
2.3505 | -82.6968
2.3508 | -82.6916
2.3511 | -82.6865
2.3514 | -82.685
2.3517 | -82.684
2.352 | -82.6823
2.3523 | -82.6791
2.3526 | -82.6726
2.3529 | -82.6743
2.3533 | -82.6776
2.3536 | -82.6771
2.3539 | -82.676
2.3542 | -82.6747
2.3545 | -82.6739
2.3548 | -82.6727
2.3551 | -82.6704
2.3554 | -82.6701
2.3557 | -82.6695
2.3561 | -82.6676
2.3564 | -82.668
2.3567 | -82.6679
2.357 | -82.6674
2.3573 | -82.6661
2.3576 | -82.6631
2.3579 | -82.6616
2.3582 | -82.6605
2.3585 | -82.6597
2.3588 | -82.6584
2.3592 | -82.6576
2.3595 | -82.6589
2.3598 | -82.6588
2.3601 | -82.6567
2.3604 | -82.6556
2.3607 | -82.6551
2.361 | -82.6549
2.3613 | -82.6535
2.3616 | -82.6529
2.362 | -82.6533
2.3623 | -82.6527
2.3626 | -82.6535
2.3629 | -82.653
2.3632 | -82.6524
2.3635 | -82.6502
2.3638 | -82.6491
2.3641 | -82.6489
2.3644 | -82.6492
2.3647 | -82.648
2.3651 | -82.648
2.3654 | -82.648
2.3657 | -82.6463
2.366 | -82.6464
2.3663 | -82.6466
2.3666 | -82.6479
2.3669 | -82.6483
2.3672 | -82.6471
2.3675 | -82.6462
2.3679 | -82.6468
2.3682 | -82.6459
2.3685 | -82.6459
2.3688 | -82.6457
2.3691 | -82.6458
2.3694 | -82.6459
2.3697 | -82.646
2.37 | -82.6458
2.3703 | -82.6461
2.3706 | -82.6474
2.371 | -82.6469
2.3713 | -82.6466
2.3716 | -82.6474
2.3719 | -82.6488
2.3722 | -82.6486
2.3725 | -82.6496
2.3728 | -82.6497
2.3731 | -82.6505
2.3734 | -82.6511
2.3738 | -82.6516
2.3741 | -82.6521
2.3744 | -82.6518
2.3747 | -82.6533
2.375 | -82.6519
2.3753 | -82.6519
2.3756 | -82.6525
2.3759 | -82.6534
2.3762 | -82.6531
2.3765 | -82.6516
2.3769 | -82.6512
2.3772 | -82.6517
2.3775 | -82.6547
2.3778 | -82.6555
2.3781 | -82.6568
2.3784 | -82.6589
2.3787 | -82.6612
2.379 | -82.662
2.3793 | -82.6621
2.3797 | -82.662
2.38 | -82.662
2.3803 | -82.6641
2.3806 | -82.6646
2.3809 | -82.6657
2.3812 | -82.6666
2.3815 | -82.6677
2.3818 | -82.6694
2.3821 | -82.672
2.3824 | -82.674
2.3828 | -82.6757
2.3831 | -82.6769
2.3834 | -82.677
2.3837 | -82.6781
2.384 | -82.6794
2.3843 | -82.6811
2.3846 | -82.6821
2.3849 | -82.6838
2.3852 | -82.6846
2.3856 | -82.6859
2.3859 | -82.6887
2.3862 | -82.6918
2.3865 | -82.6928
2.3868 | -82.6946
2.3871 | -82.6961
2.3874 | -82.6961
2.3877 | -82.6949
2.388 | -82.6953
2.3883 | -82.697
2.3887 | -82.6985
2.389 | -82.6972
2.3893 | -82.6979
2.3896 | -82.6994
2.3899 | -82.7001
2.3902 | -82.7015
2.3905 | -82.7043
2.3908 | -82.7042
2.3911 | -82.7038
2.3915 | -82.7051
2.3918 | -82.7073
2.3921 | -82.7086
2.3924 | -82.7085
2.3927 | -82.7099
2.393 | -82.7131
2.3933 | -82.7142
2.3936 | -82.7162
2.3939 | -82.7183
2.3942 | -82.7182
2.3946 | -82.7178
2.3949 | -82.7189
2.3952 | -82.7202
2.3955 | -82.7216
2.3958 | -82.7223
2.3961 | -82.7238
2.3964 | -82.7242
2.3967 | -82.7267
2.397 | -82.7298
2.3974 | -82.7311
2.3977 | -82.7331
2.398 | -82.7345
2.3983 | -82.7363
2.3986 | -82.7377
2.3989 | -82.7393
2.3992 | -82.7408
2.3995 | -82.7421
2.3998 | -82.744
2.4001 | -82.7441
2.4005 | -82.7456
2.4008 | -82.7472
2.4011 | -82.7492
2.4014 | -82.7465
2.4017 | -82.7479
2.402 | -82.7516
2.4023 | -82.7526
2.4026 | -82.7534
2.4029 | -82.7544
2.4033 | -82.7555
2.4036 | -82.7574
2.4039 | -82.7572
2.4042 | -82.7583
2.4045 | -82.7618
2.4048 | -82.7605
2.4051 | -82.7571
2.4054 | -82.7576
2.4057 | -82.759
2.406 | -82.7603
2.4064 | -82.7638
2.4067 | -82.7671
2.407 | -82.7714
2.4073 | -82.7738
2.4076 | -82.7744
2.4079 | -82.7746
2.4082 | -82.7767
2.4085 | -82.7788
2.4088 | -82.7801
2.4092 | -82.7797
2.4095 | -82.7809
2.4098 | -82.7822
2.4101 | -82.7833
2.4104 | -82.7842
2.4107 | -82.7864
2.411 | -82.7879
2.4113 | -82.7894
2.4116 | -82.7913
2.4119 | -82.7934
2.4123 | -82.7944
2.4126 | -82.7955
2.4129 | -82.7982
2.4132 | -82.8004
2.4135 | -82.8013
2.4138 | -82.8013
2.4141 | -82.8029
2.4144 | -82.8033
2.4147 | -82.8031
2.4151 | -82.8027
2.4154 | -82.8014
2.4157 | -82.8025
2.416 | -82.8054
2.4163 | -82.8067
2.4166 | -82.8053
2.4169 | -82.8058
2.4172 | -82.8065
2.4175 | -82.8076
2.4178 | -82.8097
2.4182 | -82.8105
2.4185 | -82.8132
2.4188 | -82.8137
2.4191 | -82.8153
2.4194 | -82.8144
2.4197 | -82.816
2.42 | -82.8179
2.4203 | -82.8199
2.4206 | -82.8259
2.421 | -82.834
2.4213 | -82.8352
2.4216 | -82.8359
2.4219 | -82.839
2.4222 | -82.8444
2.4225 | -82.8466
2.4228 | -82.8468
2.4231 | -82.8473
2.4234 | -82.8479
2.4237 | -82.8489
2.4241 | -82.8496
2.4244 | -82.8461
2.4247 | -82.8405
2.425 | -82.8365
2.4253 | -82.8373
2.4256 | -82.8357
2.4259 | -82.8319
2.4262 | -82.8309
2.4265 | -82.831
2.4269 | -82.8315
2.4272 | -82.8321
2.4275 | -82.8335
2.4278 | -82.833
2.4281 | -82.833
2.4284 | -82.833
2.4287 | -82.833
2.429 | -82.8335
2.4293 | -82.8338
2.4296 | -82.8333
2.43 | -82.8337
2.4303 | -82.8324
2.4306 | -82.8322
2.4309 | -82.8327
2.4312 | -82.8329
2.4315 | -82.8331
2.4318 | -82.8317
2.4321 | -82.8314
2.4324 | -82.8317
2.4328 | -82.8314
2.4331 | -82.8318
2.4334 | -82.8327
2.4337 | -82.8321
2.434 | -82.8306
2.4343 | -82.8306
2.4346 | -82.8287
2.4349 | -82.8277
2.4352 | -82.8232
2.4355 | -82.8191
2.4359 | -82.8185
2.4362 | -82.8184
2.4365 | -82.8174
2.4368 | -82.8173
2.4371 | -82.8191
2.4374 | -82.8182
2.4377 | -82.8172
2.438 | -82.8166
2.4383 | -82.8168
2.4387 | -82.816
2.439 | -82.8151
2.4393 | -82.8179
2.4396 | -82.8174
2.4399 | -82.8134
2.4402 | -82.8113
2.4405 | -82.8106
2.4408 | -82.8091
2.4411 | -82.8092
2.4414 | -82.8105
2.4418 | -82.8105
2.4421 | -82.8104
2.4424 | -82.8069
2.4427 | -82.8042
2.443 | -82.805
2.4433 | -82.8046
2.4436 | -82.8035
2.4439 | -82.8044
2.4442 | -82.8064
2.4446 | -82.8065
2.4449 | -82.8053
2.4452 | -82.8045
2.4455 | -82.806
2.4458 | -82.8065
2.4461 | -82.807
2.4464 | -82.8059
2.4467 | -82.8041
2.447 | -82.8038
2.4473 | -82.804
2.4477 | -82.8033
2.448 | -82.803
2.4483 | -82.8021
2.4486 | -82.8016
2.4489 | -82.8007
2.4492 | -82.7991
2.4495 | -82.7978
2.4498 | -82.797
2.4501 | -82.7961
2.4505 | -82.7949
2.4508 | -82.7931
2.4511 | -82.793
2.4514 | -82.7939
2.4517 | -82.7946
2.452 | -82.7933
2.4523 | -82.7893
2.4526 | -82.7889
2.4529 | -82.789
2.4532 | -82.7882
2.4536 | -82.7872
2.4539 | -82.7855
2.4542 | -82.7846
2.4545 | -82.7837
2.4548 | -82.7833
2.4551 | -82.7837
2.4554 | -82.7849
2.4557 | -82.7862
2.456 | -82.7846
2.4564 | -82.7817
2.4567 | -82.7808
2.457 | -82.7813
2.4573 | -82.7817
2.4576 | -82.7801
2.4579 | -82.7784
2.4582 | -82.7778
2.4585 | -82.7771
2.4588 | -82.7764
2.4591 | -82.7763
2.4595 | -82.7749
2.4598 | -82.7738
2.4601 | -82.7732
2.4604 | -82.7718
2.4607 | -82.7703
2.461 | -82.7688
2.4613 | -82.7672
2.4616 | -82.7668
2.4619 | -82.7668
2.4623 | -82.7661
2.4626 | -82.7644
2.4629 | -82.764
2.4632 | -82.7629
2.4635 | -82.7633
2.4638 | -82.762
2.4641 | -82.7642
2.4644 | -82.7644
2.4647 | -82.763
2.465 | -82.7619
2.4654 | -82.7608
2.4657 | -82.7598
2.466 | -82.756
2.4663 | -82.7517
2.4666 | -82.7504
2.4669 | -82.7502
2.4672 | -82.7506
2.4675 | -82.7543
2.4678 | -82.7565
2.4682 | -82.7573
2.4685 | -82.7546
2.4688 | -82.7526
2.4691 | -82.7515
2.4694 | -82.7472
2.4697 | -82.7435
2.47 | -82.742
2.4703 | -82.7435
2.4706 | -82.744
2.4709 | -82.7434
2.4713 | -82.7415
2.4716 | -82.7358
2.4719 | -82.7345
2.4722 | -82.7329
2.4725 | -82.7281
2.4728 | -82.7233
2.4731 | -82.7203
2.4734 | -82.7179
2.4737 | -82.7166
2.4741 | -82.7152
2.4744 | -82.712
2.4747 | -82.707
2.475 | -82.7039
2.4753 | -82.7021
2.4756 | -82.6992
2.4759 | -82.6975
2.4762 | -82.6966
2.4765 | -82.6951
2.4768 | -82.6954
2.4772 | -82.6939
2.4775 | -82.6924
2.4778 | -82.6904
2.4781 | -82.6898
2.4784 | -82.6884
2.4787 | -82.6848
2.479 | -82.682
2.4793 | -82.6802
2.4796 | -82.6783
2.48 | -82.6763
2.4803 | -82.675
2.4806 | -82.6734
2.4809 | -82.6713
2.4812 | -82.6676
2.4815 | -82.6638
2.4818 | -82.6627
2.4821 | -82.6616
2.4824 | -82.6599
2.4827 | -82.6587
2.4831 | -82.6572
2.4834 | -82.6544
2.4837 | -82.6533
2.484 | -82.6511
2.4843 | -82.6492
2.4846 | -82.6474
2.4849 | -82.6462
2.4852 | -82.645
2.4855 | -82.6427
2.4859 | -82.6404
2.4862 | -82.6355
2.4865 | -82.6291
2.4868 | -82.6241
2.4871 | -82.6219
2.4874 | -82.6209
2.4877 | -82.6191
2.488 | -82.617
2.4883 | -82.6149
2.4886 | -82.6118
2.489 | -82.6095
2.4893 | -82.6081
2.4896 | -82.6061
2.4899 | -82.6037
2.4902 | -82.6018
2.4905 | -82.5997
2.4908 | -82.5969
2.4911 | -82.5942
2.4914 | -82.5927
2.4918 | -82.5888
2.4921 | -82.5857
2.4924 | -82.5837
2.4927 | -82.5827
2.493 | -82.5807
2.4933 | -82.5797
2.4936 | -82.5785
2.4939 | -82.5756
2.4942 | -82.572
2.4945 | -82.5697
2.4949 | -82.5686
2.4952 | -82.5663
2.4955 | -82.5639
2.4958 | -82.5626
2.4961 | -82.561
2.4964 | -82.5583
2.4967 | -82.5569
2.497 | -82.5551
2.4973 | -82.554
2.4977 | -82.5517
2.498 | -82.5488
2.4983 | -82.5478
2.4986 | -82.5467
2.4989 | -82.5449
2.4992 | -82.5419
2.4995 | -82.54
2.4998 | -82.5376
2.5001 | -82.5361
2.5004 | -82.5345
2.5008 | -82.5322
2.5011 | -82.531
2.5014 | -82.5308
2.5017 | -82.529
2.502 | -82.5231
2.5023 | -82.5175
2.5026 | -82.5148
2.5029 | -82.5125
2.5032 | -82.5078
2.5036 | -82.5044
2.5039 | -82.5046
2.5042 | -82.5017
2.5045 | -82.4979
2.5048 | -82.4942
2.5051 | -82.4924
2.5054 | -82.4907
2.5057 | -82.4888
2.506 | -82.4874
2.5063 | -82.4862
2.5067 | -82.484
2.507 | -82.4815
2.5073 | -82.4803
2.5076 | -82.4789
2.5079 | -82.4792
2.5082 | -82.4773
2.5085 | -82.4765
2.5088 | -82.4746
2.5091 | -82.4725
2.5095 | -82.4717
2.5098 | -82.4694
2.5101 | -82.4672
2.5104 | -82.4656
2.5107 | -82.4617
2.511 | -82.4569
2.5113 | -82.4544
2.5116 | -82.4532
2.5119 | -82.4528
2.5123 | -82.4511
2.5126 | -82.4493
2.5129 | -82.4487
2.5132 | -82.4477
2.5135 | -82.4469
2.5138 | -82.4464
2.5141 | -82.4465
2.5144 | -82.4453
2.5147 | -82.4432
2.515 | -82.4418
2.5154 | -82.4405
2.5157 | -82.4392
2.516 | -82.4392
2.5163 | -82.4382
2.5166 | -82.4373
2.5169 | -82.4363
2.5172 | -82.4362
2.5175 | -82.4365
2.5178 | -82.4359
2.5182 | -82.4346
2.5185 | -82.433
2.5188 | -82.4316
2.5191 | -82.4305
2.5194 | -82.4291
2.5197 | -82.4285
2.52 | -82.4279
2.5203 | -82.4303
2.5206 | -82.4345
2.5209 | -82.4345
2.5213 | -82.4327
2.5216 | -82.4319
2.5219 | -82.4337
2.5222 | -82.4374
2.5225 | -82.4372
2.5228 | -82.4361
2.5231 | -82.435
2.5234 | -82.4329
2.5237 | -82.432
2.5241 | -82.4312
2.5244 | -82.4305
2.5247 | -82.4309
2.525 | -82.431
2.5253 | -82.429
2.5256 | -82.4265
2.5259 | -82.4265
2.5262 | -82.4251
2.5265 | -82.4254
2.5268 | -82.4245
2.5272 | -82.4222
2.5275 | -82.4206
2.5278 | -82.4223
2.5281 | -82.4226
2.5284 | -82.4219
2.5287 | -82.4212
2.529 | -82.4191
2.5293 | -82.4187
2.5296 | -82.4145
2.53 | -82.4136
2.5303 | -82.4135
2.5306 | -82.4127
2.5309 | -82.4107
2.5312 | -82.4097
2.5315 | -82.4095
2.5318 | -82.4087
2.5321 | -82.4071
2.5324 | -82.4061
2.5327 | -82.4058
2.5331 | -82.4052
2.5334 | -82.4043
2.5337 | -82.4051
2.534 | -82.4049
2.5343 | -82.403
2.5346 | -82.4024
2.5349 | -82.4012
2.5352 | -82.4
2.5355 | -82.3992
2.5359 | -82.399
2.5362 | -82.4007
2.5365 | -82.4026
2.5368 | -82.4013
2.5371 | -82.3989
2.5374 | -82.3983
2.5377 | -82.3975
2.538 | -82.3984
2.5383 | -82.3993
2.5386 | -82.4003
2.539 | -82.4014
2.5393 | -82.401
2.5396 | -82.4001
2.5399 | -82.3994
2.5402 | -82.3994
2.5405 | -82.3992
2.5408 | -82.3992
2.5411 | -82.3995
2.5414 | -82.3994
2.5418 | -82.4002
2.5421 | -82.4001
2.5424 | -82.3999
2.5427 | -82.3988
2.543 | -82.3972
2.5433 | -82.3963
2.5436 | -82.3966
2.5439 | -82.3969
2.5442 | -82.3971
2.5445 | -82.3996
2.5449 | -82.4005
2.5452 | -82.4006
2.5455 | -82.4013
2.5458 | -82.4028
2.5461 | -82.4024
2.5464 | -82.4044
2.5467 | -82.4059
2.547 | -82.4052
2.5473 | -82.4046
2.5477 | -82.4059
2.548 | -82.4074
2.5483 | -82.4076
2.5486 | -82.4088
2.5489 | -82.4095
2.5492 | -82.4089
2.5495 | -82.408
2.5498 | -82.4082
2.5501 | -82.4085
2.5504 | -82.4076
2.5508 | -82.4082
2.5511 | -82.4071
2.5514 | -82.4071
2.5517 | -82.4072
2.552 | -82.4081
2.5523 | -82.4079
2.5526 | -82.4075
2.5529 | -82.4083
2.5532 | -82.4064
2.5536 | -82.4041
2.5539 | -82.4044
2.5542 | -82.405
2.5545 | -82.4053
2.5548 | -82.4041
2.5551 | -82.4022
2.5554 | -82.4032
2.5557 | -82.4034
2.556 | -82.404
2.5563 | -82.4038
2.5567 | -82.4044
2.557 | -82.4058
2.5573 | -82.4052
2.5576 | -82.4042
2.5579 | -82.4036
2.5582 | -82.4037
2.5585 | -82.4034
2.5588 | -82.404
2.5591 | -82.405
2.5595 | -82.4072
2.5598 | -82.4101
2.5601 | -82.4119
2.5604 | -82.4122
2.5607 | -82.4131
2.561 | -82.4162
2.5613 | -82.4179
2.5616 | -82.4159
2.5619 | -82.4111
2.5622 | -82.4104
2.5626 | -82.4102
2.5629 | -82.4097
2.5632 | -82.4119
2.5635 | -82.4135
2.5638 | -82.4138
2.5641 | -82.4142
2.5644 | -82.4142
2.5647 | -82.4155
2.565 | -82.4164
2.5654 | -82.4177
2.5657 | -82.417
2.566 | -82.4164
2.5663 | -82.4171
2.5666 | -82.418
2.5669 | -82.4201
2.5672 | -82.421
2.5675 | -82.4205
2.5678 | -82.4206
2.5681 | -82.4193
2.5685 | -82.4168
2.5688 | -82.4161
2.5691 | -82.4176
2.5694 | -82.4179
2.5697 | -82.4178
2.57 | -82.4189
2.5703 | -82.4162
2.5706 | -82.4147
2.5709 | -82.4151
2.5713 | -82.4138
2.5716 | -82.4136
2.5719 | -82.414
2.5722 | -82.4147
2.5725 | -82.4158
2.5728 | -82.4178
2.5731 | -82.4193
2.5734 | -82.4191
2.5737 | -82.419
2.574 | -82.416
2.5744 | -82.4168
2.5747 | -82.4211
2.575 | -82.4253
2.5753 | -82.4258
2.5756 | -82.4262
2.5759 | -82.4252
2.5762 | -82.4231
2.5765 | -82.423
2.5768 | -82.4203
2.5772 | -82.4201
2.5775 | -82.4198
2.5778 | -82.4199
2.5781 | -82.4221
2.5784 | -82.426
2.5787 | -82.4266
2.579 | -82.4266
2.5793 | -82.4268
2.5796 | -82.4269
2.5799 | -82.4276
2.5803 | -82.4261
2.5806 | -82.4258
2.5809 | -82.4248
2.5812 | -82.4251
2.5815 | -82.4263
2.5818 | -82.4263
2.5821 | -82.4253
2.5824 | -82.4246
2.5827 | -82.4248
2.5831 | -82.4242
2.5834 | -82.424
2.5837 | -82.4229
2.584 | -82.4228
2.5843 | -82.4226
2.5846 | -82.422
2.5849 | -82.4221
2.5852 | -82.4218
2.5855 | -82.4217
2.5858 | -82.4216
2.5862 | -82.4217
2.5865 | -82.4224
2.5868 | -82.4208
2.5871 | -82.4147
2.5874 | -82.4133
2.5877 | -82.4122
2.588 | -82.411
2.5883 | -82.4109
2.5886 | -82.4124
2.589 | -82.4132
2.5893 | -82.4104
2.5896 | -82.4116
2.5899 | -82.4112
2.5902 | -82.4107
2.5905 | -82.4101
2.5908 | -82.4083
2.5911 | -82.4073
2.5914 | -82.4061
2.5917 | -82.4033
2.5921 | -82.4022
2.5924 | -82.4011
2.5927 | -82.3993
2.593 | -82.3991
2.5933 | -82.3981
2.5936 | -82.395
2.5939 | -82.395
2.5942 | -82.394
2.5945 | -82.3923
2.5949 | -82.3915
2.5952 | -82.3899
2.5955 | -82.3896
2.5958 | -82.3885
2.5961 | -82.3868
2.5964 | -82.3863
2.5967 | -82.386
2.597 | -82.3855
2.5973 | -82.383
2.5976 | -82.3814
2.598 | -82.3803
2.5983 | -82.3795
2.5986 | -82.3784
2.5989 | -82.3765
2.5992 | -82.3781
2.5995 | -82.3777
2.5998 | -82.3783
2.6001 | -82.3767
2.6004 | -82.3769
2.6008 | -82.3785
2.6011 | -82.3781
2.6014 | -82.3768
2.6017 | -82.3742
2.602 | -82.3727
2.6023 | -82.371
2.6026 | -82.3657
2.6029 | -82.361
2.6032 | -82.3583
2.6035 | -82.3564
2.6039 | -82.3545
2.6042 | -82.3531
2.6045 | -82.3506
2.6048 | -82.3479
2.6051 | -82.3447
2.6054 | -82.3439
2.6057 | -82.3419
2.606 | -82.3391
2.6063 | -82.3365
2.6067 | -82.3275
2.607 | -82.3238
2.6073 | -82.3219
2.6076 | -82.319
2.6079 | -82.3138
2.6082 | -82.3114
2.6085 | -82.3093
2.6088 | -82.3072
2.6091 | -82.3064
2.6094 | -82.3068
2.6098 | -82.3053
2.6101 | -82.3043
2.6104 | -82.3036
2.6107 | -82.304
2.611 | -82.3039
2.6113 | -82.301
2.6116 | -82.2962
2.6119 | -82.2942
2.6122 | -82.2909
2.6126 | -82.2885
2.6129 | -82.2856
2.6132 | -82.2829
2.6135 | -82.2798
2.6138 | -82.276
2.6141 | -82.2716
2.6144 | -82.2693
2.6147 | -82.2666
2.615 | -82.263
2.6153 | -82.2613
2.6157 | -82.2614
2.616 | -82.2557
2.6163 | -82.2535
2.6166 | -82.2509
2.6169 | -82.2489
2.6172 | -82.246
2.6175 | -82.2427
2.6178 | -82.2381
2.6181 | -82.2331
2.6185 | -82.2304
2.6188 | -82.2287
2.6191 | -82.2234
2.6194 | -82.2207
2.6197 | -82.218
2.62 | -82.2149
2.6203 | -82.2115
2.6206 | -82.2069
2.6209 | -82.2022
2.6212 | -82.1991
2.6216 | -82.1958
2.6219 | -82.1925
2.6222 | -82.1881
2.6225 | -82.1826
2.6228 | -82.1799
2.6231 | -82.1779
2.6234 | -82.1756
2.6237 | -82.172
2.624 | -82.1688
2.6244 | -82.1636
2.6247 | -82.1597
2.625 | -82.1558
2.6253 | -82.1521
2.6256 | -82.1491
2.6259 | -82.1464
2.6262 | -82.1423
2.6265 | -82.1375
2.6268 | -82.1312
2.6271 | -82.1232
2.6275 | -82.1196
2.6278 | -82.1162
2.6281 | -82.1125
2.6284 | -82.109
2.6287 | -82.1051
2.629 | -82.1014
2.6293 | -82.0974
2.6296 | -82.094
2.6299 | -82.0908
2.6303 | -82.0874
2.6306 | -82.0836
2.6309 | -82.08
2.6312 | -82.0753
2.6315 | -82.071
2.6318 | -82.0679
2.6321 | -82.0642
2.6324 | -82.0608
2.6327 | -82.0572
2.633 | -82.0511
2.6334 | -82.0465
2.6337 | -82.0426
2.634 | -82.037
2.6343 | -82.0328
2.6346 | -82.0281
2.6349 | -82.0233
2.6352 | -82.019
2.6355 | -82.0144
2.6358 | -82.0111
2.6362 | -82.0072
2.6365 | -81.9987
2.6368 | -81.9881
2.6371 | -81.9812
2.6374 | -81.9773
2.6377 | -81.9719
2.638 | -81.9661
2.6383 | -81.9629
2.6386 | -81.9584
2.6389 | -81.9535
2.6393 | -81.9503
2.6396 | -81.9462
2.6399 | -81.9427
2.6402 | -81.9387
2.6405 | -81.9336
2.6408 | -81.9285
2.6411 | -81.9234
2.6414 | -81.9198
2.6417 | -81.9162
2.6421 | -81.9121
2.6424 | -81.9085
2.6427 | -81.9047
2.643 | -81.902
2.6433 | -81.8982
2.6436 | -81.8943
2.6439 | -81.8921
2.6442 | -81.8895
2.6445 | -81.8851
2.6448 | -81.8809
2.6452 | -81.877
2.6455 | -81.8721
2.6458 | -81.8657
2.6461 | -81.8617
2.6464 | -81.8592
2.6467 | -81.8568
2.647 | -81.8516
2.6473 | -81.8485
2.6476 | -81.8465
2.648 | -81.8424
2.6483 | -81.8382
2.6486 | -81.8342
2.6489 | -81.8295
2.6492 | -81.8262
2.6495 | -81.8249
2.6498 | -81.8213
2.6501 | -81.8158
2.6504 | -81.8127
2.6508 | -81.8092
2.6511 | -81.806
2.6514 | -81.7996
2.6517 | -81.7954
2.652 | -81.7909
2.6523 | -81.7844
2.6526 | -81.7789
2.6529 | -81.7755
2.6532 | -81.773
2.6535 | -81.7645
2.6539 | -81.76
2.6542 | -81.7536
2.6545 | -81.7486
2.6548 | -81.7441
2.6551 | -81.7403
2.6554 | -81.7367
2.6557 | -81.7313
2.656 | -81.7272
2.6563 | -81.7218
2.6567 | -81.7146
2.657 | -81.7091
2.6573 | -81.7048
2.6576 | -81.7012
2.6579 | -81.6965
2.6582 | -81.6915
2.6585 | -81.6875
2.6588 | -81.6812
2.6591 | -81.6748
2.6594 | -81.6697
2.6598 | -81.6649
2.6601 | -81.6594
2.6604 | -81.6557
2.6607 | -81.6516
2.661 | -81.6473
2.6613 | -81.6428
2.6616 | -81.6363
2.6619 | -81.6315
2.6622 | -81.6264
2.6626 | -81.6231
2.6629 | -81.6187
2.6632 | -81.6146
2.6635 | -81.6106
2.6638 | -81.6054
2.6641 | -81.601
2.6644 | -81.5976
2.6647 | -81.5944
2.665 | -81.5909
2.6653 | -81.589
2.6657 | -81.5879
2.666 | -81.5856
2.6663 | -81.581
2.6666 | -81.5755
2.6669 | -81.5715
2.6672 | -81.5667
2.6675 | -81.5614
2.6678 | -81.5585
2.6681 | -81.5553
2.6685 | -81.55
2.6688 | -81.5454
2.6691 | -81.5418
2.6694 | -81.5384
2.6697 | -81.535
2.67 | -81.5309
2.6703 | -81.5274
2.6706 | -81.524
2.6709 | -81.5152
2.6712 | -81.5059
2.6716 | -81.5036
2.6719 | -81.5003
2.6722 | -81.4953
2.6725 | -81.4911
2.6728 | -81.4876
2.6731 | -81.4832
2.6734 | -81.4791
2.6737 | -81.4755
2.674 | -81.4714
2.6744 | -81.4653
2.6747 | -81.4595
2.675 | -81.456
2.6753 | -81.4526
2.6756 | -81.4485
2.6759 | -81.4449
2.6762 | -81.4417
2.6765 | -81.4376
2.6768 | -81.433
2.6771 | -81.4283
2.6775 | -81.4267
2.6778 | -81.4242
2.6781 | -81.4191
2.6784 | -81.4148
2.6787 | -81.4105
2.679 | -81.405
2.6793 | -81.4023
2.6796 | -81.3971
2.6799 | -81.3914
2.6803 | -81.385
2.6806 | -81.3796
2.6809 | -81.3759
2.6812 | -81.3727
2.6815 | -81.368
2.6818 | -81.3625
2.6821 | -81.3589
2.6824 | -81.3581
2.6827 | -81.3565
2.683 | -81.3536
2.6834 | -81.3505
2.6837 | -81.348
2.684 | -81.3457
2.6843 | -81.3408
2.6846 | -81.3384
2.6849 | -81.3355
2.6852 | -81.3327
2.6855 | -81.3294
2.6858 | -81.3256
2.6862 | -81.3221
2.6865 | -81.3175
2.6868 | -81.3126
2.6871 | -81.31
2.6874 | -81.3069
2.6877 | -81.3044
2.688 | -81.2988
2.6883 | -81.2948
2.6886 | -81.2913
2.6889 | -81.2879
2.6893 | -81.2844
2.6896 | -81.2805
2.6899 | -81.2764
2.6902 | -81.2728
2.6905 | -81.2684
2.6908 | -81.2654
2.6911 | -81.2625
2.6914 | -81.2591
2.6917 | -81.256
2.6921 | -81.2548
2.6924 | -81.2513
2.6927 | -81.2455
2.693 | -81.2423
2.6933 | -81.2391
2.6936 | -81.2357
2.6939 | -81.2314
2.6942 | -81.2276
2.6945 | -81.2217
2.6948 | -81.2167
2.6952 | -81.2132
2.6955 | -81.2092
2.6958 | -81.204
2.6961 | -81.2
2.6964 | -81.1959
2.6967 | -81.1935
2.697 | -81.19
2.6973 | -81.1847
2.6976 | -81.1802
2.698 | -81.1765
2.6983 | -81.1729
2.6986 | -81.1677
2.6989 | -81.1625
2.6992 | -81.1594
2.6995 | -81.1561
2.6998 | -81.1528
2.7001 | -81.1488
2.7004 | -81.1445
2.7007 | -81.1416
2.7011 | -81.1382
2.7014 | -81.135
2.7017 | -81.1317
2.702 | -81.127
2.7023 | -81.1232
2.7026 | -81.1189
2.7029 | -81.1141
2.7032 | -81.1104
2.7035 | -81.1075
2.7039 | -81.1028
2.7042 | -81.095
2.7045 | -81.0903
2.7048 | -81.0874
2.7051 | -81.0847
2.7054 | -81.0809
2.7057 | -81.0771
2.706 | -81.0741
2.7063 | -81.0674
2.7066 | -81.0628
2.707 | -81.0602
2.7073 | -81.0593
2.7076 | -81.0578
2.7079 | -81.0535
2.7082 | -81.0484
2.7085 | -81.0447
2.7088 | -81.0414
2.7091 | -81.0385
2.7094 | -81.0367
2.7098 | -81.0347
2.7101 | -81.0314
2.7104 | -81.0263
2.7107 | -81.0224
2.711 | -81.0198
2.7113 | -81.0181
2.7116 | -81.0151
2.7119 | -81.0142
2.7122 | -81.0111
2.7125 | -81.0079
2.7129 | -81.0045
2.7132 | -80.998
2.7135 | -80.9931
2.7138 | -80.9893
2.7141 | -80.9861
2.7144 | -80.9841
2.7147 | -80.9827
2.715 | -80.9802
2.7153 | -80.9775
2.7157 | -80.9754
2.716 | -80.9729
2.7163 | -80.9703
2.7166 | -80.9676
2.7169 | -80.9657
2.7172 | -80.9632
2.7175 | -80.9615
2.7178 | -80.9592
2.7181 | -80.9575
2.7184 | -80.9555
2.7188 | -80.953
2.7191 | -80.9504
2.7194 | -80.9479
2.7197 | -80.9449
2.72 | -80.9409
2.7203 | -80.9374
2.7206 | -80.9351
2.7209 | -80.9317
2.7212 | -80.9292
2.7216 | -80.9262
2.7219 | -80.9215
2.7222 | -80.9157
2.7225 | -80.911
2.7228 | -80.9089
2.7231 | -80.9064
2.7234 | -80.9039
2.7237 | -80.9023
2.724 | -80.8987
2.7243 | -80.8949
2.7247 | -80.8926
2.725 | -80.8876
2.7253 | -80.8841
2.7256 | -80.882
2.7259 | -80.8802
2.7262 | -80.8793
2.7265 | -80.8782
2.7268 | -80.8758
2.7271 | -80.8735
2.7275 | -80.8713
2.7278 | -80.8673
2.7281 | -80.8607
2.7284 | -80.8539
2.7287 | -80.8494
2.729 | -80.8462
2.7293 | -80.8444
2.7296 | -80.8418
2.7299 | -80.8391
2.7302 | -80.8369
2.7306 | -80.8355
2.7309 | -80.8336
2.7312 | -80.8305
2.7315 | -80.8276
2.7318 | -80.8261
2.7321 | -80.823
2.7324 | -80.8218
2.7327 | -80.8191
2.733 | -80.8138
2.7334 | -80.8106
2.7337 | -80.8077
2.734 | -80.8005
2.7343 | -80.795
2.7346 | -80.7901
2.7349 | -80.7886
2.7352 | -80.7865
2.7355 | -80.7837
2.7358 | -80.7814
2.7361 | -80.7791
2.7365 | -80.7771
2.7368 | -80.7747
2.7371 | -80.7728
2.7374 | -80.7714
2.7377 | -80.7701
2.738 | -80.7671
2.7383 | -80.7606
2.7386 | -80.757
2.7389 | -80.756
2.7393 | -80.7553
2.7396 | -80.7529
2.7399 | -80.7494
2.7402 | -80.7466
2.7405 | -80.7434
2.7408 | -80.7396
2.7411 | -80.7357
2.7414 | -80.7318
2.7417 | -80.7302
2.742 | -80.729
2.7424 | -80.7266
2.7427 | -80.7245
2.743 | -80.7196
2.7433 | -80.7164
2.7436 | -80.7141
2.7439 | -80.712
2.7442 | -80.7095
2.7445 | -80.707
2.7448 | -80.705
2.7452 | -80.7028
2.7455 | -80.7007
2.7458 | -80.6987
2.7461 | -80.6981
2.7464 | -80.6982
2.7467 | -80.6958
2.747 | -80.6937
2.7473 | -80.6915
2.7476 | -80.6903
2.7479 | -80.6883
2.7483 | -80.6849
2.7486 | -80.682
2.7489 | -80.6787
2.7492 | -80.6769
2.7495 | -80.6749
2.7498 | -80.6734
2.7501 | -80.671
2.7504 | -80.6666
2.7507 | -80.6648
2.7511 | -80.6635
2.7514 | -80.6599
2.7517 | -80.6565
2.752 | -80.6543
2.7523 | -80.6531
2.7526 | -80.6505
2.7529 | -80.6488
2.7532 | -80.6476
2.7535 | -80.6469
2.7538 | -80.645
2.7542 | -80.6412
2.7545 | -80.6374
2.7548 | -80.6357
2.7551 | -80.6325
2.7554 | -80.6304
2.7557 | -80.6294
2.756 | -80.6268
2.7563 | -80.6243
2.7566 | -80.6228
2.757 | -80.6206
2.7573 | -80.6183
2.7576 | -80.6147
2.7579 | -80.6121
2.7582 | -80.6098
2.7585 | -80.6086
2.7588 | -80.6059
2.7591 | -80.6042
2.7594 | -80.6025
2.7597 | -80.6025
2.7601 | -80.6001
2.7604 | -80.5977
2.7607 | -80.5962
2.761 | -80.5939
2.7613 | -80.5919
2.7616 | -80.5899
2.7619 | -80.5878
2.7622 | -80.5863
2.7625 | -80.5852
2.7629 | -80.5835
2.7632 | -80.5818
2.7635 | -80.5783
2.7638 | -80.576
2.7641 | -80.5733
2.7644 | -80.5697
2.7647 | -80.5671
2.765 | -80.564
2.7653 | -80.5589
2.7656 | -80.555
2.766 | -80.553
2.7663 | -80.5513
2.7666 | -80.5491
2.7669 | -80.5464
2.7672 | -80.5443
2.7675 | -80.5423
2.7678 | -80.5403
2.7681 | -80.5387
2.7684 | -80.5376
2.7688 | -80.5358
2.7691 | -80.5337
2.7694 | -80.531
2.7697 | -80.5292
2.77 | -80.5279
2.7703 | -80.5262
2.7706 | -80.5235
2.7709 | -80.52
2.7712 | -80.5181
2.7715 | -80.5166
2.7719 | -80.5145
2.7722 | -80.5126
2.7725 | -80.5099
2.7728 | -80.5071
2.7731 | -80.5056
2.7734 | -80.5063
2.7737 | -80.501
2.774 | -80.499
2.7743 | -80.4971
2.7747 | -80.4947
2.775 | -80.4913
2.7753 | -80.4878
2.7756 | -80.4862
2.7759 | -80.4851
2.7762 | -80.4826
2.7765 | -80.4801
2.7768 | -80.4779
2.7771 | -80.4751
2.7774 | -80.474
2.7778 | -80.4741
2.7781 | -80.4723
2.7784 | -80.469
2.7787 | -80.4681
2.779 | -80.4663
2.7793 | -80.4635
2.7796 | -80.4623
2.7799 | -80.458
2.7802 | -80.4554
2.7806 | -80.4534
2.7809 | -80.449
2.7812 | -80.4467
2.7815 | -80.4429
2.7818 | -80.4392
2.7821 | -80.4371
2.7824 | -80.435
2.7827 | -80.4329
2.783 | -80.4295
2.7833 | -80.4269
2.7837 | -80.4228
2.784 | -80.4206
2.7843 | -80.4192
2.7846 | -80.4175
2.7849 | -80.4131
2.7852 | -80.4089
2.7855 | -80.4064
2.7858 | -80.403
2.7861 | -80.4002
2.7865 | -80.3977
2.7868 | -80.3963
2.7871 | -80.3968
2.7874 | -80.395
2.7877 | -80.3914
2.788 | -80.3893
2.7883 | -80.384
2.7886 | -80.3815
2.7889 | -80.3791
2.7893 | -80.376
2.7896 | -80.3735
2.7899 | -80.3714
2.7902 | -80.3708
2.7905 | -80.371
2.7908 | -80.3684
2.7911 | -80.3631
2.7914 | -80.3581
2.7917 | -80.3547
2.792 | -80.3526
2.7924 | -80.3498
2.7927 | -80.3457
2.793 | -80.3412
2.7933 | -80.3362
2.7936 | -80.3315
2.7939 | -80.3279
2.7942 | -80.3283
2.7945 | -80.3278
2.7948 | -80.3252
2.7952 | -80.3211
2.7955 | -80.3184
2.7958 | -80.3167
2.7961 | -80.3157
2.7964 | -80.3135
2.7967 | -80.3109
2.797 | -80.3086
2.7973 | -80.3062
2.7976 | -80.3033
2.7979 | -80.3004
2.7983 | -80.2975
2.7986 | -80.2946
2.7989 | -80.2923
2.7992 | -80.2899
2.7995 | -80.2877
2.7998 | -80.286
2.8001 | -80.2835
2.8004 | -80.2801
2.8007 | -80.2777
2.8011 | -80.2753
2.8014 | -80.2721
2.8017 | -80.27
2.802 | -80.2683
2.8023 | -80.2665
2.8026 | -80.2642
2.8029 | -80.2624
2.8032 | -80.2601
2.8035 | -80.2577
2.8038 | -80.2549
2.8042 | -80.2528
2.8045 | -80.2501
2.8048 | -80.2477
2.8051 | -80.2437
2.8054 | -80.2402
2.8057 | -80.2368
2.806 | -80.2334
2.8063 | -80.2309
2.8066 | -80.2287
2.807 | -80.2248
2.8073 | -80.223
2.8076 | -80.2212
2.8079 | -80.2199
2.8082 | -80.2155
2.8085 | -80.2114
2.8088 | -80.2087
2.8091 | -80.2063
2.8094 | -80.2038
2.8097 | -80.202
2.8101 | -80.1974
2.8104 | -80.1926
2.8107 | -80.1895
2.811 | -80.1874
2.8113 | -80.1844
2.8116 | -80.18
2.8119 | -80.1767
2.8122 | -80.1739
2.8125 | -80.1698
2.8129 | -80.1671
2.8132 | -80.1634
2.8135 | -80.1593
2.8138 | -80.1558
2.8141 | -80.1526
2.8144 | -80.1497
2.8147 | -80.1469
2.815 | -80.1439
2.8153 | -80.1401
2.8156 | -80.1376
2.816 | -80.1363
2.8163 | -80.135
2.8166 | -80.1355
2.8169 | -80.133
2.8172 | -80.1305
2.8175 | -80.1282
2.8178 | -80.1252
2.8181 | -80.122
2.8184 | -80.1209
2.8188 | -80.119
2.8191 | -80.1165
2.8194 | -80.1137
2.8197 | -80.1113
2.82 | -80.1086
2.8203 | -80.1057
2.8206 | -80.1028
2.8209 | -80.099
2.8212 | -80.0954
2.8215 | -80.0916
2.8219 | -80.0893
2.8222 | -80.0881
2.8225 | -80.0856
2.8228 | -80.0843
2.8231 | -80.082
2.8234 | -80.0761
2.8237 | -80.0718
2.824 | -80.0698
2.8243 | -80.0672
2.8247 | -80.0647
2.825 | -80.0621
2.8253 | -80.0601
2.8256 | -80.0578
2.8259 | -80.0526
2.8262 | -80.0475
2.8265 | -80.0473
2.8268 | -80.0486
2.8271 | -80.0476
2.8274 | -80.0441
2.8278 | -80.0423
2.8281 | -80.0412
2.8284 | -80.0378
2.8287 | -80.0322
2.829 | -80.0279
2.8293 | -80.0236
2.8296 | -80.0182
2.8299 | -80.0141
2.8302 | -80.0109
2.8306 | -80.0091
2.8309 | -80.0073
2.8312 | -80.0053
2.8315 | -80.0005
2.8318 | -79.9958
2.8321 | -79.9925
2.8324 | -79.9893
2.8327 | -79.9865
2.833 | -79.9834
2.8333 | -79.9807
2.8337 | -79.9779
2.834 | -79.9754
2.8343 | -79.9731
2.8346 | -79.97
2.8349 | -79.9666
2.8352 | -79.9626
2.8355 | -79.96
2.8358 | -79.9563
2.8361 | -79.9514
2.8365 | -79.9487
2.8368 | -79.9457
2.8371 | -79.9437
2.8374 | -79.9408
2.8377 | -79.9377
2.838 | -79.9347
2.8383 | -79.9313
2.8386 | -79.9285
2.8389 | -79.9258
2.8392 | -79.9227
2.8396 | -79.9196
2.8399 | -79.9161
2.8402 | -79.9141
2.8405 | -79.9113
2.8408 | -79.9079
2.8411 | -79.9017
2.8414 | -79.8958
2.8417 | -79.8935
2.842 | -79.8911
2.8424 | -79.8904
2.8427 | -79.8876
2.843 | -79.8844
2.8433 | -79.8824
2.8436 | -79.8803
2.8439 | -79.8775
2.8442 | -79.875
2.8445 | -79.8721
2.8448 | -79.8711
2.8451 | -79.8673
2.8455 | -79.8659
2.8458 | -79.8635
2.8461 | -79.8604
2.8464 | -79.8563
2.8467 | -79.8524
2.847 | -79.8495
2.8473 | -79.8447
2.8476 | -79.8411
2.8479 | -79.8396
2.8483 | -79.838
2.8486 | -79.836
2.8489 | -79.8328
2.8492 | -79.8298
2.8495 | -79.8273
2.8498 | -79.825
2.8501 | -79.8221
2.8504 | -79.8191
2.8507 | -79.8166
2.851 | -79.8141
2.8514 | -79.8125
2.8517 | -79.8111
2.852 | -79.8077
2.8523 | -79.8046
2.8526 | -79.8016
2.8529 | -79.7991
2.8532 | -79.7964
2.8535 | -79.795
2.8538 | -79.7926
2.8542 | -79.7904
2.8545 | -79.7877
2.8548 | -79.7848
2.8551 | -79.781
2.8554 | -79.7771
2.8557 | -79.7753
2.856 | -79.7709
2.8563 | -79.7672
2.8566 | -79.7625
2.8569 | -79.7583
2.8573 | -79.7539
2.8576 | -79.7508
2.8579 | -79.7428
2.8582 | -79.7401
2.8585 | -79.7368
2.8588 | -79.7333
2.8591 | -79.7308
2.8594 | -79.7273
2.8597 | -79.7249
2.8601 | -79.7215
2.8604 | -79.7186
2.8607 | -79.7166
2.861 | -79.7141
2.8613 | -79.7111
2.8616 | -79.7083
2.8619 | -79.7046
2.8622 | -79.7015
2.8625 | -79.6992
2.8628 | -79.6966
2.8632 | -79.6931
2.8635 | -79.692
2.8638 | -79.6924
2.8641 | -79.6899
2.8644 | -79.6875
2.8647 | -79.6854
2.865 | -79.6823
2.8653 | -79.6791
2.8656 | -79.6761
2.866 | -79.6725
2.8663 | -79.67
2.8666 | -79.6681
2.8669 | -79.667
2.8672 | -79.6659
2.8675 | -79.6642
2.8678 | -79.6613
2.8681 | -79.6577
2.8684 | -79.6546
2.8687 | -79.6511
2.8691 | -79.6482
2.8694 | -79.6457
2.8697 | -79.6417
2.87 | -79.6387
2.8703 | -79.6346
2.8706 | -79.6303
2.8709 | -79.6284
2.8712 | -79.6263
2.8715 | -79.6238
2.8719 | -79.622
2.8722 | -79.6203
2.8725 | -79.6179
2.8728 | -79.616
2.8731 | -79.6141
2.8734 | -79.6125
2.8737 | -79.6095
2.874 | -79.6047
2.8743 | -79.6014
2.8746 | -79.5977
2.875 | -79.5958
2.8753 | -79.5937
2.8756 | -79.5911
2.8759 | -79.5903
2.8762 | -79.5881
2.8765 | -79.5856
2.8768 | -79.5826
2.8771 | -79.5802
2.8774 | -79.5774
2.8778 | -79.5759
2.8781 | -79.573
2.8784 | -79.5708
2.8787 | -79.5688
2.879 | -79.566
2.8793 | -79.5628
2.8796 | -79.5578
2.8799 | -79.5521
2.8802 | -79.5507
2.8805 | -79.5481
2.8809 | -79.546
2.8812 | -79.543
2.8815 | -79.5412
2.8818 | -79.5403
2.8821 | -79.5378
2.8824 | -79.5345
2.8827 | -79.5327
2.883 | -79.5304
2.8833 | -79.5288
2.8837 | -79.5273
2.884 | -79.5243
2.8843 | -79.5216
2.8846 | -79.519
2.8849 | -79.5185
2.8852 | -79.5165
2.8855 | -79.5144
2.8858 | -79.5132
2.8861 | -79.5103
2.8864 | -79.5082
2.8868 | -79.5079
2.8871 | -79.506
2.8874 | -79.5042
2.8877 | -79.5014
2.888 | -79.5
2.8883 | -79.4986
2.8886 | -79.4983
2.8889 | -79.4976
2.8892 | -79.4948
2.8896 | -79.4929
2.8899 | -79.4925
2.8902 | -79.49
2.8905 | -79.4888
2.8908 | -79.4861
2.8911 | -79.4833
2.8914 | -79.4806
2.8917 | -79.4777
2.892 | -79.4755
2.8923 | -79.4737
2.8927 | -79.4717
2.893 | -79.4714
2.8933 | -79.4699
2.8936 | -79.4674
2.8939 | -79.4666
2.8942 | -79.463
2.8945 | -79.4579
2.8948 | -79.4548
2.8951 | -79.4542
2.8955 | -79.4504
2.8958 | -79.4474
2.8961 | -79.4461
2.8964 | -79.4446
2.8967 | -79.4442
2.897 | -79.4431
2.8973 | -79.4419
2.8976 | -79.4393
2.8979 | -79.4365
2.8982 | -79.4351
2.8986 | -79.4343
2.8989 | -79.4358
2.8992 | -79.435
2.8995 | -79.4308
2.8998 | -79.426
2.9001 | -79.4246
2.9004 | -79.4229
2.9007 | -79.422
2.901 | -79.4205
2.9014 | -79.4199
2.9017 | -79.418
2.902 | -79.4159
2.9023 | -79.4134
2.9026 | -79.4107
2.9029 | -79.4095
2.9032 | -79.4072
2.9035 | -79.4057
2.9038 | -79.4028
2.9041 | -79.4
2.9045 | -79.3964
2.9048 | -79.3945
2.9051 | -79.3915
2.9054 | -79.3891
2.9057 | -79.3889
2.906 | -79.3879
2.9063 | -79.3862
2.9066 | -79.385
2.9069 | -79.3831
2.9073 | -79.3819
2.9076 | -79.3809
2.9079 | -79.3801
2.9082 | -79.3787
2.9085 | -79.3793
2.9088 | -79.3795
2.9091 | -79.3778
2.9094 | -79.3757
2.9097 | -79.3742
2.9101 | -79.3718
2.9104 | -79.3689
2.9107 | -79.3675
2.911 | -79.3661
2.9113 | -79.3641
2.9116 | -79.3631
2.9119 | -79.3629
2.9122 | -79.3613
2.9125 | -79.3604
2.9128 | -79.3592
2.9132 | -79.3576
2.9135 | -79.356
2.9138 | -79.3543
2.9141 | -79.353
2.9144 | -79.3511
2.9147 | -79.3494
2.915 | -79.3485
2.9153 | -79.3475
2.9156 | -79.3473
2.916 | -79.3479
2.9163 | -79.3462
2.9166 | -79.343
2.9169 | -79.3421
2.9172 | -79.3408
2.9175 | -79.3381
2.9178 | -79.3375
2.9181 | -79.3385
2.9184 | -79.3387
2.9187 | -79.3379
2.9191 | -79.3368
2.9194 | -79.3356
2.9197 | -79.3348
2.92 | -79.3342
2.9203 | -79.3341
2.9206 | -79.3346
2.9209 | -79.335
2.9212 | -79.3337
2.9215 | -79.3338
2.9219 | -79.334
2.9222 | -79.3333
2.9225 | -79.333
2.9228 | -79.3322
2.9231 | -79.3315
2.9234 | -79.3311
2.9237 | -79.3305
2.924 | -79.3296
2.9243 | -79.3294
2.9246 | -79.3316
2.925 | -79.3336
2.9253 | -79.3346
2.9256 | -79.3351
2.9259 | -79.3337
2.9262 | -79.3315
2.9265 | -79.3283
2.9268 | -79.3269
2.9271 | -79.3266
2.9274 | -79.326
2.9278 | -79.3259
2.9281 | -79.326
2.9284 | -79.323
2.9287 | -79.3228
2.929 | -79.3226
2.9293 | -79.322
2.9296 | -79.3222
2.9299 | -79.3221
2.9302 | -79.3239
2.9305 | -79.3225
2.9309 | -79.3187
2.9312 | -79.3196
2.9315 | -79.3218
2.9318 | -79.3223
2.9321 | -79.3202
2.9324 | -79.3195
2.9327 | -79.319
2.933 | -79.3181
2.9333 | -79.3164
2.9337 | -79.3152
2.934 | -79.3141
2.9343 | -79.3133
2.9346 | -79.3132
2.9349 | -79.3127
2.9352 | -79.3121
2.9355 | -79.3145
2.9358 | -79.3146
2.9361 | -79.3135
2.9364 | -79.3134
2.9368 | -79.3125
2.9371 | -79.3129
2.9374 | -79.3123
2.9377 | -79.3126
2.938 | -79.3133
2.9383 | -79.3138
2.9386 | -79.3141
2.9389 | -79.3145
2.9392 | -79.3141
2.9396 | -79.3125
2.9399 | -79.3111
2.9402 | -79.31
2.9405 | -79.3088
2.9408 | -79.3096
2.9411 | -79.3083
2.9414 | -79.3074
2.9417 | -79.3068
2.942 | -79.3057
2.9423 | -79.3045
2.9427 | -79.3027
2.943 | -79.3016
2.9433 | -79.3009
2.9436 | -79.3003
2.9439 | -79.2995
2.9442 | -79.2993
2.9445 | -79.2996
2.9448 | -79.2965
2.9451 | -79.2958
2.9455 | -79.295
2.9458 | -79.2942
2.9461 | -79.2945
2.9464 | -79.2941
2.9467 | -79.2908
2.947 | -79.2851
2.9473 | -79.2791
2.9476 | -79.2781
2.9479 | -79.2772
2.9482 | -79.2764
2.9486 | -79.2766
2.9489 | -79.2762
2.9492 | -79.2744
2.9495 | -79.2706
2.9498 | -79.2702
2.9501 | -79.2697
2.9504 | -79.2668
2.9507 | -79.2652
2.951 | -79.2656
2.9514 | -79.2643
2.9517 | -79.2637
2.952 | -79.2636
2.9523 | -79.2632
2.9526 | -79.2631
2.9529 | -79.2628
2.9532 | -79.2615
2.9535 | -79.2596
2.9538 | -79.2597
2.9541 | -79.2584
2.9545 | -79.2573
2.9548 | -79.2563
2.9551 | -79.2559
2.9554 | -79.2554
2.9557 | -79.2546
2.956 | -79.2547
2.9563 | -79.2527
2.9566 | -79.2501
2.9569 | -79.2459
2.9573 | -79.2443
2.9576 | -79.2429
2.9579 | -79.2397
2.9582 | -79.239
2.9585 | -79.2375
2.9588 | -79.2338
2.9591 | -79.2325
2.9594 | -79.2317
2.9597 | -79.2305
2.96 | -79.2292
2.9604 | -79.2278
2.9607 | -79.2264
2.961 | -79.2259
2.9613 | -79.2244
2.9616 | -79.2219
2.9619 | -79.2209
2.9622 | -79.2203
2.9625 | -79.2203
2.9628 | -79.2196
2.9632 | -79.2178
2.9635 | -79.2172
2.9638 | -79.2167
2.9641 | -79.2159
2.9644 | -79.2163
2.9647 | -79.2162
2.965 | -79.2146
2.9653 | -79.2139
2.9656 | -79.2112
2.9659 | -79.2102
2.9663 | -79.2097
2.9666 | -79.2079
2.9669 | -79.2061
2.9672 | -79.2041
2.9675 | -79.2028
2.9678 | -79.2018
2.9681 | -79.2009
2.9684 | -79.2
2.9687 | -79.1989
2.9691 | -79.1974
2.9694 | -79.1948
2.9697 | -79.1938
2.97 | -79.1923
2.9703 | -79.1909
2.9706 | -79.191
2.9709 | -79.19
2.9712 | -79.1891
2.9715 | -79.1882
2.9718 | -79.187
2.9722 | -79.1856
2.9725 | -79.1835
2.9728 | -79.1821
2.9731 | -79.1817
2.9734 | -79.1806
2.9737 | -79.1772
2.974 | -79.175
2.9743 | -79.1739
2.9746 | -79.1732
2.975 | -79.1728
2.9753 | -79.1726
2.9756 | -79.1706
2.9759 | -79.1688
2.9762 | -79.1677
2.9765 | -79.166
2.9768 | -79.1647
2.9771 | -79.1646
2.9774 | -79.1637
2.9777 | -79.1621
2.9781 | -79.1598
2.9784 | -79.1583
2.9787 | -79.1582
2.979 | -79.1569
2.9793 | -79.1558
2.9796 | -79.1556
2.9799 | -79.1532
2.9802 | -79.1534
2.9805 | -79.1528
2.9809 | -79.1527
2.9812 | -79.1536
2.9815 | -79.1523
2.9818 | -79.1525
2.9821 | -79.1522
2.9824 | -79.1511
2.9827 | -79.15
2.983 | -79.1493
2.9833 | -79.1482
2.9836 | -79.1481
2.984 | -79.1467
2.9843 | -79.1463
2.9846 | -79.1452
2.9849 | -79.1443
2.9852 | -79.1431
2.9855 | -79.1432
2.9858 | -79.1433
2.9861 | -79.1425
2.9864 | -79.1415
2.9868 | -79.1406
2.9871 | -79.1394
2.9874 | -79.1391
2.9877 | -79.1375
2.988 | -79.1359
2.9883 | -79.1342
2.9886 | -79.1339
2.9889 | -79.133
2.9892 | -79.1333
2.9895 | -79.1322
2.9899 | -79.1313
2.9902 | -79.1313
2.9905 | -79.1303
2.9908 | -79.1288
2.9911 | -79.1285
2.9914 | -79.1268
2.9917 | -79.1265
2.992 | -79.1274
2.9923 | -79.1256
2.9927 | -79.1265
2.993 | -79.1258
2.9933 | -79.125
2.9936 | -79.1248
2.9939 | -79.1245
2.9942 | -79.1248
2.9945 | -79.1258
2.9948 | -79.126
2.9951 | -79.1256
2.9954 | -79.1254
2.9958 | -79.1243
2.9961 | -79.1233
2.9964 | -79.1232
2.9967 | -79.123
2.997 | -79.1216
2.9973 | -79.1195
2.9976 | -79.1185
2.9979 | -79.1172
2.9982 | -79.1153
2.9986 | -79.1141
2.9989 | -79.1129
2.9992 | -79.1121
2.9995 | -79.1117
2.9998 | -79.1104
3.0001 | -79.1092
3.0004 | -79.1084
3.0007 | -79.1081
3.001 | -79.1075
3.0013 | -79.1078
3.0017 | -79.1079
3.002 | -79.107
3.0023 | -79.1062
3.0026 | -79.1055
3.0029 | -79.1051
3.0032 | -79.1052
3.0035 | -79.1049
3.0038 | -79.1046
3.0041 | -79.1042
3.0045 | -79.1033
3.0048 | -79.1026
3.0051 | -79.1019
3.0054 | -79.1015
3.0057 | -79.1011
3.006 | -79.0999
3.0063 | -79.0991
3.0066 | -79.0989
3.0069 | -79.0995
3.0072 | -79.1004
3.0076 | -79.1005
3.0079 | -79.1007
3.0082 | -79.1014
3.0085 | -79.1013
3.0088 | -79.1013
3.0091 | -79.103
3.0094 | -79.1024
3.0097 | -79.1016
3.01 | -79.1014
3.0104 | -79.1032
3.0107 | -79.1045
3.011 | -79.1053
3.0113 | -79.1046
3.0116 | -79.1046
3.0119 | -79.1053
3.0122 | -79.1056
3.0125 | -79.1067
3.0128 | -79.1072
3.0131 | -79.107
3.0135 | -79.1071
3.0138 | -79.1082
3.0141 | -79.1093
3.0144 | -79.1098
3.0147 | -79.1093
3.015 | -79.1104
3.0153 | -79.1107
3.0156 | -79.1107
3.0159 | -79.1118
3.0163 | -79.1113
3.0166 | -79.1114
3.0169 | -79.1118
3.0172 | -79.1122
3.0175 | -79.1128
3.0178 | -79.1133
3.0181 | -79.1144
3.0184 | -79.1145
3.0187 | -79.1143
3.019 | -79.1152
3.0194 | -79.1177
3.0197 | -79.1178
3.02 | -79.1186
3.0203 | -79.1186
3.0206 | -79.1179
3.0209 | -79.118
3.0212 | -79.1186
3.0215 | -79.1191
3.0218 | -79.1197
3.0222 | -79.1199
3.0225 | -79.1206
3.0228 | -79.1215
3.0231 | -79.1224
3.0234 | -79.1228
3.0237 | -79.1245
3.024 | -79.1265
3.0243 | -79.1278
3.0246 | -79.1293
3.0249 | -79.1298
3.0253 | -79.1297
3.0256 | -79.1303
3.0259 | -79.1312
3.0262 | -79.1304
3.0265 | -79.1306
3.0268 | -79.1342
3.0271 | -79.1369
3.0274 | -79.1377
3.0277 | -79.1391
3.0281 | -79.1408
3.0284 | -79.1448
3.0287 | -79.1468
3.029 | -79.1478
3.0293 | -79.1492
3.0296 | -79.1518
3.0299 | -79.155
3.0302 | -79.1566
3.0305 | -79.1577
3.0308 | -79.1595
3.0312 | -79.1615
3.0315 | -79.1624
3.0318 | -79.1629
3.0321 | -79.1647
3.0324 | -79.1689
3.0327 | -79.1723
3.033 | -79.1751
3.0333 | -79.1786
3.0336 | -79.181
3.034 | -79.1821
3.0343 | -79.1832
3.0346 | -79.1847
3.0349 | -79.1861
3.0352 | -79.1873
3.0355 | -79.1907
3.0358 | -79.1943
3.0361 | -79.1962
3.0364 | -79.1976
3.0367 | -79.1984
3.0371 | -79.2025
3.0374 | -79.2038
3.0377 | -79.2046
3.038 | -79.2066
3.0383 | -79.2087
3.0386 | -79.211
3.0389 | -79.2138
3.0392 | -79.2162
3.0395 | -79.2196
3.0399 | -79.2211
3.0402 | -79.224
3.0405 | -79.2282
3.0408 | -79.2313
3.0411 | -79.2338
3.0414 | -79.2361
3.0417 | -79.2387
3.042 | -79.2422
3.0423 | -79.2458
3.0426 | -79.2488
3.043 | -79.2515
3.0433 | -79.254
3.0436 | -79.2564
3.0439 | -79.2595
3.0442 | -79.263
3.0445 | -79.2655
3.0448 | -79.2685
3.0451 | -79.2711
3.0454 | -79.2738
3.0458 | -79.2803
3.0461 | -79.2861
3.0464 | -79.292
3.0467 | -79.2942
3.047 | -79.2971
3.0473 | -79.2999
3.0476 | -79.3026
3.0479 | -79.3061
3.0482 | -79.3091
3.0485 | -79.3125
3.0489 | -79.3164
3.0492 | -79.3191
3.0495 | -79.3218
3.0498 | -79.3259
3.0501 | -79.3301
3.0504 | -79.3337
3.0507 | -79.3373
3.051 | -79.3412
3.0513 | -79.3461
3.0517 | -79.3499
3.052 | -79.3538
3.0523 | -79.3588
3.0526 | -79.3637
3.0529 | -79.3677
3.0532 | -79.3711
3.0535 | -79.3744
3.0538 | -79.3785
3.0541 | -79.3829
3.0545 | -79.3859
3.0548 | -79.3898
3.0551 | -79.3966
3.0554 | -79.3999
3.0557 | -79.4029
3.056 | -79.4082
3.0563 | -79.4131
3.0566 | -79.4169
3.0569 | -79.4195
3.0572 | -79.4238
3.0576 | -79.4284
3.0579 | -79.4328
3.0582 | -79.4374
3.0585 | -79.4416
3.0588 | -79.4457
3.0591 | -79.4494
3.0594 | -79.4536
3.0597 | -79.4581
3.06 | -79.4618
3.0604 | -79.4664
3.0607 | -79.4721
3.061 | -79.4777
3.0613 | -79.4824
3.0616 | -79.4885
3.0619 | -79.4938
3.0622 | -79.5
3.0625 | -79.506
3.0628 | -79.5125
3.0631 | -79.5179
3.0635 | -79.5258
3.0638 | -79.5324
3.0641 | -79.5383
3.0644 | -79.5431
3.0647 | -79.5489
3.065 | -79.5542
3.0653 | -79.5584
3.0656 | -79.5646
3.0659 | -79.5697
3.0663 | -79.5752
3.0666 | -79.5804
3.0669 | -79.5857
3.0672 | -79.5911
3.0675 | -79.5965
3.0678 | -79.6017
3.0681 | -79.6069
3.0684 | -79.6122
3.0687 | -79.6172
3.069 | -79.6222
3.0694 | -79.6274
3.0697 | -79.6336
3.07 | -79.6385
3.0703 | -79.644
3.0706 | -79.6495
3.0709 | -79.6541
3.0712 | -79.66
3.0715 | -79.6658
3.0718 | -79.6717
3.0722 | -79.6775
3.0725 | -79.6829
3.0728 | -79.6886
3.0731 | -79.6962
3.0734 | -79.7015
3.0737 | -79.7073
3.074 | -79.7131
3.0743 | -79.7193
3.0746 | -79.7256
3.0749 | -79.732
3.0753 | -79.7382
3.0756 | -79.745
3.0759 | -79.7493
3.0762 | -79.7534
3.0765 | -79.7609
3.0768 | -79.7679
3.0771 | -79.7752
3.0774 | -79.7824
3.0777 | -79.788
3.0781 | -79.7918
3.0784 | -79.7965
3.0787 | -79.8026
3.079 | -79.8088
3.0793 | -79.8142
3.0796 | -79.8213
3.0799 | -79.8281
3.0802 | -79.8354
3.0805 | -79.8433
3.0808 | -79.8515
3.0812 | -79.8585
3.0815 | -79.8665
3.0818 | -79.8733
3.0821 | -79.8805
3.0824 | -79.8861
3.0827 | -79.8915
3.083 | -79.8977
3.0833 | -79.9033
3.0836 | -79.9094
3.084 | -79.9156
3.0843 | -79.9223
3.0846 | -79.9293
3.0849 | -79.9361
3.0852 | -79.9426
3.0855 | -79.9497
3.0858 | -79.9565
3.0861 | -79.9632
3.0864 | -79.9701
3.0867 | -79.9775
3.0871 | -79.987
3.0874 | -79.9979
3.0877 | -80.0049
3.088 | -80.012
3.0883 | -80.0185
3.0886 | -80.0247
3.0889 | -80.0302
3.0892 | -80.0368
3.0895 | -80.0439
3.0899 | -80.0516
3.0902 | -80.0581
3.0905 | -80.0652
3.0908 | -80.0719
3.0911 | -80.0787
3.0914 | -80.0854
3.0917 | -80.0938
3.092 | -80.1017
3.0923 | -80.1089
3.0926 | -80.1155
3.093 | -80.122
3.0933 | -80.1288
3.0936 | -80.1344
3.0939 | -80.1401
3.0942 | -80.1483
3.0945 | -80.1559
3.0948 | -80.1624
3.0951 | -80.1689
3.0954 | -80.1756
3.0958 | -80.1839
3.0961 | -80.1908
3.0964 | -80.1978
3.0967 | -80.2051
3.097 | -80.2118
3.0973 | -80.2196
3.0976 | -80.2273
3.0979 | -80.2337
3.0982 | -80.2424
3.0985 | -80.2494
3.0989 | -80.2581
3.0992 | -80.2664
3.0995 | -80.2745
3.0998 | -80.2812
3.1001 | -80.2897
3.1004 | -80.2978
3.1007 | -80.3049
3.101 | -80.312
3.1013 | -80.319
3.1017 | -80.3271
3.102 | -80.3353
3.1023 | -80.3431
3.1026 | -80.3497
3.1029 | -80.3588
3.1032 | -80.3689
3.1035 | -80.3778
3.1038 | -80.3845
3.1041 | -80.3933
3.1044 | -80.4019
3.1048 | -80.4086
3.1051 | -80.415
3.1054 | -80.4213
3.1057 | -80.4279
3.106 | -80.4348
3.1063 | -80.4433
3.1066 | -80.4511
3.1069 | -80.4576
3.1072 | -80.4648
3.1076 | -80.4715
3.1079 | -80.4788
3.1082 | -80.4885
3.1085 | -80.4984
3.1088 | -80.5039
3.1091 | -80.5109
3.1094 | -80.5186
3.1097 | -80.5251
3.11 | -80.5314
3.1103 | -80.5398
3.1107 | -80.5481
3.111 | -80.555
3.1113 | -80.5612
3.1116 | -80.5683
3.1119 | -80.5766
3.1122 | -80.5837
3.1125 | -80.5921
3.1128 | -80.6007
3.1131 | -80.61
3.1135 | -80.6184
3.1138 | -80.6274
3.1141 | -80.6342
3.1144 | -80.6424
3.1147 | -80.65
3.115 | -80.6559
3.1153 | -80.6624
3.1156 | -80.6693
3.1159 | -80.6757
3.1162 | -80.6831
3.1166 | -80.6899
3.1169 | -80.6966
3.1172 | -80.703
3.1175 | -80.7102
3.1178 | -80.7167
3.1181 | -80.7242
3.1184 | -80.7308
3.1187 | -80.7386
3.119 | -80.7466
3.1194 | -80.7532
3.1197 | -80.7602
3.12 | -80.768
3.1203 | -80.774
3.1206 | -80.7796
3.1209 | -80.7869
3.1212 | -80.7932
3.1215 | -80.8
3.1218 | -80.8068
3.1221 | -80.814
3.1225 | -80.8201
3.1228 | -80.8267
3.1231 | -80.8333
3.1234 | -80.8403
3.1237 | -80.8477
3.124 | -80.8545
3.1243 | -80.8622
3.1246 | -80.8693
3.1249 | -80.8756
3.1253 | -80.8817
3.1256 | -80.8879
3.1259 | -80.8943
3.1262 | -80.9028
3.1265 | -80.9121
3.1268 | -80.917
3.1271 | -80.9238
3.1274 | -80.9306
3.1277 | -80.9398
3.128 | -80.947
3.1284 | -80.9535
3.1287 | -80.9595
3.129 | -80.964
3.1293 | -80.9698
3.1296 | -80.9756
3.1299 | -80.981
3.1302 | -80.9871
3.1305 | -80.9935
3.1308 | -80.999
3.1312 | -81.0057
3.1315 | -81.0119
3.1318 | -81.0188
3.1321 | -81.0252
3.1324 | -81.0305
3.1327 | -81.0362
3.133 | -81.0415
3.1333 | -81.0472
3.1336 | -81.0533
3.1339 | -81.0593
3.1343 | -81.0653
3.1346 | -81.0718
3.1349 | -81.079
3.1352 | -81.0853
3.1355 | -81.092
3.1358 | -81.0977
3.1361 | -81.1034
3.1364 | -81.109
3.1367 | -81.1161
3.1371 | -81.1252
3.1374 | -81.1354
3.1377 | -81.1418
3.138 | -81.1468
3.1383 | -81.1531
3.1386 | -81.1595
3.1389 | -81.1658
3.1392 | -81.1747
3.1395 | -81.1802
3.1398 | -81.1863
3.1402 | -81.1922
3.1405 | -81.1986
3.1408 | -81.2054
3.1411 | -81.2114
3.1414 | -81.2176
3.1417 | -81.2265
3.142 | -81.2377
3.1423 | -81.246
3.1426 | -81.2518
3.143 | -81.2587
3.1433 | -81.2673
3.1436 | -81.2738
3.1439 | -81.2798
3.1442 | -81.2857
3.1445 | -81.2934
3.1448 | -81.2991
3.1451 | -81.3051
3.1454 | -81.3111
3.1457 | -81.317
3.1461 | -81.3213
3.1464 | -81.3274
3.1467 | -81.334
3.147 | -81.3397
3.1473 | -81.3455
3.1476 | -81.3518
3.1479 | -81.3563
3.1482 | -81.3611
3.1485 | -81.3661
3.1489 | -81.3727
3.1492 | -81.3786
3.1495 | -81.3842
3.1498 | -81.3896
3.1501 | -81.3947
3.1504 | -81.4015
3.1507 | -81.4076
3.151 | -81.4138
3.1513 | -81.4207
3.1516 | -81.4263
3.152 | -81.4325
3.1523 | -81.4391
3.1526 | -81.4458
3.1529 | -81.4532
3.1532 | -81.4581
3.1535 | -81.463
3.1538 | -81.4674
3.1541 | -81.4741
3.1544 | -81.4809
3.1548 | -81.4884
3.1551 | -81.4946
3.1554 | -81.5005
3.1557 | -81.5063
3.156 | -81.5126
3.1563 | -81.5188
3.1566 | -81.5241
3.1569 | -81.5332
3.1572 | -81.5391
3.1575 | -81.5451
3.1579 | -81.5513
3.1582 | -81.5568
3.1585 | -81.563
3.1588 | -81.5735
3.1591 | -81.5864
3.1594 | -81.6003
3.1597 | -81.606
3.16 | -81.6112
3.1603 | -81.6168
3.1607 | -81.622
3.161 | -81.6281
3.1613 | -81.634
3.1616 | -81.6408
3.1619 | -81.6513
3.1622 | -81.6573
3.1625 | -81.6632
3.1628 | -81.6691
3.1631 | -81.6778
3.1634 | -81.6843
3.1638 | -81.689
3.1641 | -81.6948
3.1644 | -81.6999
3.1647 | -81.7049
3.165 | -81.7096
3.1653 | -81.7151
3.1656 | -81.7206
3.1659 | -81.7266
3.1662 | -81.7345
3.1666 | -81.739
3.1669 | -81.7441
3.1672 | -81.7502
3.1675 | -81.7546
3.1678 | -81.7586
3.1681 | -81.7638
3.1684 | -81.7686
3.1687 | -81.7727
3.169 | -81.7803
3.1693 | -81.7861
3.1697 | -81.7939
3.17 | -81.8047
3.1703 | -81.8109
3.1706 | -81.8182
3.1709 | -81.8269
3.1712 | -81.8319
3.1715 | -81.837
3.1718 | -81.8425
3.1721 | -81.8513
3.1725 | -81.857
3.1728 | -81.8618
3.1731 | -81.8663
3.1734 | -81.8717
3.1737 | -81.8777
3.174 | -81.883
3.1743 | -81.8877
3.1746 | -81.8936
3.1749 | -81.9017
3.1752 | -81.907
3.1756 | -81.9115
3.1759 | -81.9165
3.1762 | -81.9211
3.1765 | -81.9272
3.1768 | -81.9317
3.1771 | -81.9354
3.1774 | -81.9398
3.1777 | -81.9444
3.178 | -81.947
3.1784 | -81.9515
3.1787 | -81.9552
3.179 | -81.961
3.1793 | -81.9678
3.1796 | -81.9728
3.1799 | -81.9768
3.1802 | -81.9832
3.1805 | -81.988
3.1808 | -81.9942
3.1811 | -81.9999
3.1815 | -82.0048
3.1818 | -82.0093
3.1821 | -82.0143
3.1824 | -82.0189
3.1827 | -82.0243
3.183 | -82.0311
3.1833 | -82.0378
3.1836 | -82.0427
3.1839 | -82.0476
3.1843 | -82.0533
3.1846 | -82.0576
3.1849 | -82.0615
3.1852 | -82.0663
3.1855 | -82.0709
3.1858 | -82.0737
3.1861 | -82.0788
3.1864 | -82.0865
3.1867 | -82.0925
3.187 | -82.097
3.1874 | -82.1066
3.1877 | -82.116
3.188 | -82.1224
3.1883 | -82.1268
3.1886 | -82.1309
3.1889 | -82.1346
3.1892 | -82.1392
3.1895 | -82.1436
3.1898 | -82.1474
3.1902 | -82.1536
3.1905 | -82.1586
3.1908 | -82.1631
3.1911 | -82.1692
3.1914 | -82.173
3.1917 | -82.1774
3.192 | -82.1821
3.1923 | -82.1845
3.1926 | -82.1897
3.193 | -82.1949
3.1933 | -82.1995
3.1936 | -82.2035
3.1939 | -82.2053
3.1942 | -82.2079
3.1945 | -82.2137
3.1948 | -82.2182
3.1951 | -82.2226
3.1954 | -82.2264
3.1957 | -82.2304
3.1961 | -82.2356
3.1964 | -82.2396
3.1967 | -82.2445
3.197 | -82.2485
3.1973 | -82.2521
3.1976 | -82.2564
3.1979 | -82.2609
3.1982 | -82.2656
3.1985 | -82.2706
3.1989 | -82.2745
3.1992 | -82.2795
3.1995 | -82.2847
3.1998 | -82.2886
3.2001 | -82.2926
3.2004 | -82.2973
3.2007 | -82.3025
3.201 | -82.3074
3.2013 | -82.3115
3.2016 | -82.3174
3.202 | -82.3228
3.2023 | -82.3266
3.2026 | -82.3308
3.2029 | -82.3353
3.2032 | -82.3398
3.2035 | -82.3456
3.2038 | -82.3507
3.2041 | -82.3549
3.2044 | -82.3598
3.2048 | -82.3643
3.2051 | -82.368
3.2054 | -82.3715
3.2057 | -82.3767
3.206 | -82.3815
3.2063 | -82.3848
3.2066 | -82.39
3.2069 | -82.3944
3.2072 | -82.3991
3.2075 | -82.4032
3.2079 | -82.4088
3.2082 | -82.4157
3.2085 | -82.419
3.2088 | -82.423
3.2091 | -82.427
3.2094 | -82.432
3.2097 | -82.4382
3.21 | -82.4395
3.2103 | -82.4423
3.2106 | -82.4475
3.211 | -82.4507
3.2113 | -82.4548
3.2116 | -82.4584
3.2119 | -82.4617
3.2122 | -82.4646
3.2125 | -82.4683
3.2128 | -82.4732
3.2131 | -82.4782
3.2134 | -82.4811
3.2138 | -82.485
3.2141 | -82.4903
3.2144 | -82.4935
3.2147 | -82.4977
3.215 | -82.502
3.2153 | -82.506
3.2156 | -82.5106
3.2159 | -82.515
3.2162 | -82.5195
3.2165 | -82.5233
3.2169 | -82.5268
3.2172 | -82.5305
3.2175 | -82.5358
3.2178 | -82.5397
3.2181 | -82.5424
3.2184 | -82.5454
3.2187 | -82.5511
3.219 | -82.5538
3.2193 | -82.5566
3.2197 | -82.5603
3.22 | -82.5636
3.2203 | -82.567
3.2206 | -82.5707
3.2209 | -82.575
3.2212 | -82.5787
3.2215 | -82.5827
3.2218 | -82.5855
3.2221 | -82.5895
3.2225 | -82.5948
3.2228 | -82.6
3.2231 | -82.604
3.2234 | -82.6073
3.2237 | -82.6109
3.224 | -82.6144
3.2243 | -82.6183
3.2246 | -82.6218
3.2249 | -82.6253
3.2252 | -82.6284
3.2256 | -82.6317
3.2259 | -82.6352
3.2262 | -82.6404
3.2265 | -82.6451
3.2268 | -82.6484
3.2271 | -82.6519
3.2274 | -82.6553
3.2277 | -82.6566
3.228 | -82.6591
3.2284 | -82.6639
3.2287 | -82.6665
3.229 | -82.6691
3.2293 | -82.6725
3.2296 | -82.6767
3.2299 | -82.6819
3.2302 | -82.6853
3.2305 | -82.6884
3.2308 | -82.6922
3.2311 | -82.6966
3.2315 | -82.6997
3.2318 | -82.7041
3.2321 | -82.7078
3.2324 | -82.71
3.2327 | -82.7141
3.233 | -82.7176
3.2333 | -82.7262
3.2336 | -82.7317
3.2339 | -82.7351
3.2343 | -82.7376
3.2346 | -82.7406
3.2349 | -82.7432
3.2352 | -82.7446
3.2355 | -82.7468
3.2358 | -82.7506
3.2361 | -82.754
3.2364 | -82.7567
3.2367 | -82.76
3.237 | -82.7623
3.2374 | -82.7644
3.2377 | -82.7663
3.238 | -82.7685
3.2383 | -82.7723
3.2386 | -82.776
3.2389 | -82.7805
3.2392 | -82.7841
3.2395 | -82.7866
3.2398 | -82.7898
3.2402 | -82.7924
3.2405 | -82.7946
3.2408 | -82.7978
3.2411 | -82.8013
3.2414 | -82.8023
3.2417 | -82.8039
3.242 | -82.806
3.2423 | -82.8088
3.2426 | -82.8121
3.2429 | -82.8167
3.2433 | -82.8188
3.2436 | -82.8217
3.2439 | -82.8247
3.2442 | -82.8322
3.2445 | -82.8383
3.2448 | -82.8402
3.2451 | -82.8423
3.2454 | -82.8453
3.2457 | -82.8475
3.2461 | -82.8483
3.2464 | -82.8489
3.2467 | -82.8509
3.247 | -82.8517
3.2473 | -82.8533
3.2476 | -82.8556
3.2479 | -82.859
3.2482 | -82.8609
3.2485 | -82.8626
3.2488 | -82.8646
3.2492 | -82.8653
3.2495 | -82.8649
3.2498 | -82.8669
3.2501 | -82.8708
3.2504 | -82.8736
3.2507 | -82.8758
3.251 | -82.8765
3.2513 | -82.8784
3.2516 | -82.8802
3.252 | -82.8815
3.2523 | -82.8829
3.2526 | -82.8865
3.2529 | -82.8883
3.2532 | -82.8901
3.2535 | -82.8902
3.2538 | -82.8931
3.2541 | -82.8959
3.2544 | -82.8967
3.2547 | -82.8992
3.2551 | -82.9027
3.2554 | -82.9069
3.2557 | -82.9096
3.256 | -82.9118
3.2563 | -82.9153
3.2566 | -82.9192
3.2569 | -82.9191
3.2572 | -82.92
3.2575 | -82.922
3.2579 | -82.9235
3.2582 | -82.927
3.2585 | -82.9293
3.2588 | -82.9318
3.2591 | -82.9326
3.2594 | -82.9351
3.2597 | -82.9378
3.26 | -82.9392
3.2603 | -82.9404
3.2606 | -82.9414
3.261 | -82.9422
3.2613 | -82.9442
3.2616 | -82.945
3.2619 | -82.946
3.2622 | -82.9474
3.2625 | -82.9493
3.2628 | -82.9518
3.2631 | -82.9529
3.2634 | -82.9547
3.2638 | -82.9567
3.2641 | -82.9586
3.2644 | -82.9616
3.2647 | -82.9654
3.265 | -82.9665
3.2653 | -82.9673
3.2656 | -82.9688
3.2659 | -82.9706
3.2662 | -82.9738
3.2665 | -82.9756
3.2669 | -82.9771
3.2672 | -82.9797
3.2675 | -82.9811
3.2678 | -82.9825
3.2681 | -82.9847
3.2684 | -82.9862
3.2687 | -82.9891
3.269 | -82.9925
3.2693 | -82.9947
3.2697 | -82.9962
3.27 | -82.9973
3.2703 | -82.9979
3.2706 | -82.9984
3.2709 | -82.9997
3.2712 | -83.0008
3.2715 | -83.0026
3.2718 | -83.0029
3.2721 | -83.0033
3.2724 | -83.0045
3.2728 | -83.0056
3.2731 | -83.0076
3.2734 | -83.0098
3.2737 | -83.0112
3.274 | -83.0136
3.2743 | -83.0162
3.2746 | -83.0174
3.2749 | -83.0178
3.2752 | -83.018
3.2756 | -83.0193
3.2759 | -83.0209
3.2762 | -83.0206
3.2765 | -83.021
3.2768 | -83.0232
3.2771 | -83.0274
3.2774 | -83.0289
3.2777 | -83.0301
3.278 | -83.0299
3.2783 | -83.0309
3.2787 | -83.0325
3.279 | -83.0335
3.2793 | -83.034
3.2796 | -83.0345
3.2799 | -83.0342
3.2802 | -83.0354
3.2805 | -83.0383
3.2808 | -83.0398
3.2811 | -83.0409
3.2815 | -83.0412
3.2818 | -83.0412
3.2821 | -83.0416
3.2824 | -83.0425
3.2827 | -83.0432
3.283 | -83.0426
3.2833 | -83.0419
3.2836 | -83.0429
3.2839 | -83.0443
3.2842 | -83.0476
3.2846 | -83.0508
3.2849 | -83.0515
3.2852 | -83.0513
3.2855 | -83.0517
3.2858 | -83.0537
3.2861 | -83.0554
3.2864 | -83.055
3.2867 | -83.0559
3.287 | -83.0572
3.2874 | -83.0585
3.2877 | -83.0598
3.288 | -83.0608
3.2883 | -83.0614
3.2886 | -83.0622
3.2889 | -83.0628
3.2892 | -83.0662
3.2895 | -83.0694
3.2898 | -83.0707
3.2901 | -83.0714
3.2905 | -83.0748
3.2908 | -83.0764
3.2911 | -83.0767
3.2914 | -83.0774
3.2917 | -83.0768
3.292 | -83.0765
3.2923 | -83.0756
3.2926 | -83.0765
3.2929 | -83.0779
3.2933 | -83.0806
3.2936 | -83.0818
3.2939 | -83.0817
3.2942 | -83.0808
3.2945 | -83.081
3.2948 | -83.0824
3.2951 | -83.0829
3.2954 | -83.084
3.2957 | -83.0849
3.296 | -83.0859
3.2964 | -83.0874
3.2967 | -83.0888
3.297 | -83.0897
3.2973 | -83.0909
3.2976 | -83.0944
3.2979 | -83.0973
3.2982 | -83.0983
3.2985 | -83.0992
3.2988 | -83.1
3.2992 | -83.1013
3.2995 | -83.1034
3.2998 | -83.1081
3.3001 | -83.1112
3.3004 | -83.1125
3.3007 | -83.1134
3.301 | -83.1146
3.3013 | -83.1154
3.3016 | -83.1162
3.3019 | -83.1169
3.3023 | -83.1174
3.3026 | -83.1179
3.3029 | -83.1187
3.3032 | -83.1187
3.3035 | -83.1194
3.3038 | -83.1201
3.3041 | -83.1214
3.3044 | -83.1231
3.3047 | -83.1233
3.3051 | -83.1235
3.3054 | -83.1257
3.3057 | -83.1254
3.306 | -83.125
3.3063 | -83.1244
3.3066 | -83.1231
3.3069 | -83.1223
3.3072 | -83.123
3.3075 | -83.1237
3.3078 | -83.1245
3.3082 | -83.1252
3.3085 | -83.1258
3.3088 | -83.1267
3.3091 | -83.1276
3.3094 | -83.1272
3.3097 | -83.128
3.31 | -83.1291
3.3103 | -83.1298
3.3106 | -83.1303
3.311 | -83.1327
3.3113 | -83.1354
3.3116 | -83.1362
3.3119 | -83.1371
3.3122 | -83.1384
3.3125 | -83.1393
3.3128 | -83.1398
3.3131 | -83.1408
3.3134 | -83.1422
3.3137 | -83.1436
3.3141 | -83.145
3.3144 | -83.1456
3.3147 | -83.1457
3.315 | -83.1467
3.3153 | -83.1477
3.3156 | -83.1491
3.3159 | -83.1508
3.3162 | -83.1504
3.3165 | -83.1533
3.3169 | -83.1542
3.3172 | -83.1535
3.3175 | -83.1544
3.3178 | -83.1556
3.3181 | -83.157
3.3184 | -83.1577
3.3187 | -83.1586
3.319 | -83.1598
3.3193 | -83.1594
3.3196 | -83.1589
3.32 | -83.1602
3.3203 | -83.1611
3.3206 | -83.1599
3.3209 | -83.1605
3.3212 | -83.1603
3.3215 | -83.1602
3.3218 | -83.1596
3.3221 | -83.1602
3.3224 | -83.1607
3.3228 | -83.1611
3.3231 | -83.1632
3.3234 | -83.166
3.3237 | -83.1673
3.324 | -83.1687
3.3243 | -83.1699
3.3246 | -83.1717
3.3249 | -83.1721
3.3252 | -83.1725
3.3255 | -83.173
3.3259 | -83.1739
3.3262 | -83.1759
3.3265 | -83.1767
3.3268 | -83.1781
3.3271 | -83.1787
3.3274 | -83.1789
3.3277 | -83.18
3.328 | -83.1805
3.3283 | -83.1813
3.3287 | -83.1822
3.329 | -83.1837
3.3293 | -83.1845
3.3296 | -83.1855
3.3299 | -83.1881
3.3302 | -83.1885
3.3305 | -83.1879
3.3308 | -83.1878
3.3311 | -83.1877
3.3314 | -83.1876
3.3318 | -83.1876
3.3321 | -83.1884
3.3324 | -83.1899
3.3327 | -83.1897
3.333 | -83.189
3.3333 | -83.19
3.3336 | -83.1899
3.3339 | -83.1901
3.3342 | -83.1909
3.3346 | -83.1912
3.3349 | -83.1905
3.3352 | -83.1903
3.3355 | -83.1932
3.3358 | -83.1947
3.3361 | -83.1942
3.3364 | -83.194
3.3367 | -83.1937
3.337 | -83.1938
3.3373 | -83.1939
3.3377 | -83.1937
3.338 | -83.1928
3.3383 | -83.1919
3.3386 | -83.1913
3.3389 | -83.1916
3.3392 | -83.193
3.3395 | -83.1932
3.3398 | -83.1902
3.3401 | -83.189
3.3405 | -83.1888
3.3408 | -83.1887
3.3411 | -83.1885
3.3414 | -83.1874
3.3417 | -83.1827
3.342 | -83.1809
3.3423 | -83.1802
3.3426 | -83.1804
3.3429 | -83.1806
3.3432 | -83.1807
3.3436 | -83.1807
3.3439 | -83.1809
3.3442 | -83.1815
3.3445 | -83.182
3.3448 | -83.1832
3.3451 | -83.1842
3.3454 | -83.1858
3.3457 | -83.1857
3.346 | -83.1851
3.3464 | -83.1849
3.3467 | -83.1847
3.347 | -83.1842
3.3473 | -83.1849
3.3476 | -83.1845
3.3479 | -83.1842
3.3482 | -83.1846
3.3485 | -83.1845
3.3488 | -83.1844
3.3491 | -83.1848
3.3495 | -83.1849
3.3498 | -83.1847
3.3501 | -83.184
3.3504 | -83.1825
3.3507 | -83.1821
3.351 | -83.1827
3.3513 | -83.1829
3.3516 | -83.1827
3.3519 | -83.1836
3.3523 | -83.1843
3.3526 | -83.1842
3.3529 | -83.184
3.3532 | -83.1836
3.3535 | -83.1844
3.3538 | -83.1831
3.3541 | -83.1828
3.3544 | -83.1843
3.3547 | -83.1842
3.355 | -83.1841
3.3554 | -83.1838
3.3557 | -83.1831
3.356 | -83.1828
3.3563 | -83.1821
3.3566 | -83.1819
3.3569 | -83.1808
3.3572 | -83.1792
3.3575 | -83.1796
3.3578 | -83.1804
3.3582 | -83.18
3.3585 | -83.1803
3.3588 | -83.1806
3.3591 | -83.1796
3.3594 | -83.1787
3.3597 | -83.1788
3.36 | -83.1787
3.3603 | -83.1777
3.3606 | -83.1773
3.3609 | -83.1761
3.3613 | -83.1733
3.3616 | -83.1718
3.3619 | -83.1712
3.3622 | -83.1706
3.3625 | -83.17
3.3628 | -83.1694
3.3631 | -83.1692
3.3634 | -83.1684
3.3637 | -83.169
3.3641 | -83.17
3.3644 | -83.1681
3.3647 | -83.1664
3.365 | -83.1657
3.3653 | -83.1655
3.3656 | -83.1663
3.3659 | -83.1656
3.3662 | -83.1646
3.3665 | -83.1638
3.3669 | -83.1628
3.3672 | -83.1619
3.3675 | -83.1605
3.3678 | -83.1588
3.3681 | -83.1584
3.3684 | -83.158
3.3687 | -83.1583
3.369 | -83.1575
3.3693 | -83.1567
3.3696 | -83.1553
3.37 | -83.1541
3.3703 | -83.154
3.3706 | -83.1532
3.3709 | -83.152
3.3712 | -83.1516
3.3715 | -83.1514
3.3718 | -83.1522
3.3721 | -83.1526
3.3724 | -83.1527
3.3728 | -83.1551
3.3731 | -83.1551
3.3734 | -83.1545
3.3737 | -83.1532
3.374 | -83.1518
3.3743 | -83.1502
3.3746 | -83.151
3.3749 | -83.1523
3.3752 | -83.1526
3.3755 | -83.151
3.3759 | -83.1498
3.3762 | -83.1487
3.3765 | -83.1485
3.3768 | -83.1482
3.3771 | -83.1476
3.3774 | -83.1467
3.3777 | -83.1472
3.378 | -83.1473
3.3783 | -83.1478
3.3787 | -83.1494
3.379 | -83.1499
3.3793 | -83.1493
3.3796 | -83.1491
3.3799 | -83.1492
3.3802 | -83.1491
3.3805 | -83.1476
3.3808 | -83.1475
3.3811 | -83.1471
3.3814 | -83.1463
3.3818 | -83.1451
3.3821 | -83.1435
3.3824 | -83.1432
3.3827 | -83.1426
3.383 | -83.1415
3.3833 | -83.1411
3.3836 | -83.1413
3.3839 | -83.1423
3.3842 | -83.1422
3.3846 | -83.1421
3.3849 | -83.1421
3.3852 | -83.142
3.3855 | -83.1414
3.3858 | -83.14
3.3861 | -83.1395
3.3864 | -83.1402
3.3867 | -83.1395
3.387 | -83.1385
3.3873 | -83.1398
3.3877 | -83.1395
3.388 | -83.1385
3.3883 | -83.1382
3.3886 | -83.1377
3.3889 | -83.1379
3.3892 | -83.138
3.3895 | -83.1372
3.3898 | -83.1361
3.3901 | -83.1351
3.3905 | -83.1345
3.3908 | -83.1344
3.3911 | -83.134
3.3914 | -83.1321
3.3917 | -83.1304
3.392 | -83.1294
3.3923 | -83.1291
3.3926 | -83.129
3.3929 | -83.1292
3.3932 | -83.1296
3.3936 | -83.1306
3.3939 | -83.13
3.3942 | -83.1284
3.3945 | -83.1288
3.3948 | -83.1295
3.3951 | -83.1295
3.3954 | -83.1307
3.3957 | -83.1317
3.396 | -83.1326
3.3964 | -83.1323
3.3967 | -83.132
3.397 | -83.1332
3.3973 | -83.1328
3.3976 | -83.132
3.3979 | -83.1326
3.3982 | -83.132
3.3985 | -83.1319
3.3988 | -83.1314
3.3991 | -83.1311
3.3995 | -83.129
3.3998 | -83.1265
3.4001 | -83.1265
3.4004 | -83.1265
3.4007 | -83.1268
3.401 | -83.127
3.4013 | -83.1262
3.4016 | -83.1265
3.4019 | -83.1272
3.4023 | -83.1272
3.4026 | -83.1283
3.4029 | -83.1255
3.4032 | -83.1214
3.4035 | -83.1209
3.4038 | -83.1209
3.4041 | -83.1206
3.4044 | -83.1194
3.4047 | -83.1177
3.405 | -83.1168
3.4054 | -83.118
3.4057 | -83.1194
3.406 | -83.1204
3.4063 | -83.1219
3.4066 | -83.1229
3.4069 | -83.126
3.4072 | -83.1273
3.4075 | -83.1275
3.4078 | -83.1281
3.4082 | -83.1279
3.4085 | -83.1278
3.4088 | -83.1285
3.4091 | -83.1283
3.4094 | -83.1286
3.4097 | -83.1285
3.41 | -83.1283
3.4103 | -83.1285
3.4106 | -83.1287
3.4109 | -83.1284
3.4113 | -83.128
3.4116 | -83.1272
3.4119 | -83.1272
3.4122 | -83.1275
3.4125 | -83.1279
3.4128 | -83.1278
3.4131 | -83.1288
3.4134 | -83.1291
3.4137 | -83.1303
3.4141 | -83.1312
3.4144 | -83.1316
3.4147 | -83.1318
3.415 | -83.1323
3.4153 | -83.135
3.4156 | -83.1369
3.4159 | -83.137
3.4162 | -83.1372
3.4165 | -83.1373
3.4168 | -83.1359
3.4172 | -83.1358
3.4175 | -83.1366
3.4178 | -83.1378
3.4181 | -83.139
3.4184 | -83.14
3.4187 | -83.1404
3.419 | -83.1406
3.4193 | -83.1412
3.4196 | -83.1424
3.42 | -83.1433
3.4203 | -83.144
3.4206 | -83.1451
3.4209 | -83.1449
3.4212 | -83.1462
3.4215 | -83.1457
3.4218 | -83.1458
3.4221 | -83.1489
3.4224 | -83.1506
3.4227 | -83.152
3.4231 | -83.1519
3.4234 | -83.1519
3.4237 | -83.1534
3.424 | -83.1567
3.4243 | -83.1589
3.4246 | -83.1596
3.4249 | -83.1605
3.4252 | -83.1609
3.4255 | -83.1613
3.4259 | -83.1618
3.4262 | -83.1622
3.4265 | -83.1635
3.4268 | -83.1639
3.4271 | -83.1652
3.4274 | -83.1667
3.4277 | -83.1674
3.428 | -83.1686
3.4283 | -83.1695
3.4286 | -83.169
3.429 | -83.1692
3.4293 | -83.1694
3.4296 | -83.1704
3.4299 | -83.1716
3.4302 | -83.1722
3.4305 | -83.172
3.4308 | -83.1728
3.4311 | -83.1731
3.4314 | -83.1751
3.4318 | -83.1745
3.4321 | -83.1739
3.4324 | -83.1742
3.4327 | -83.1753
3.433 | -83.1766
3.4333 | -83.1774
3.4336 | -83.1785
3.4339 | -83.1775
3.4342 | -83.1778
3.4345 | -83.1785
3.4349 | -83.1777
3.4352 | -83.1776
3.4355 | -83.1777
3.4358 | -83.1791
3.4361 | -83.1805
3.4364 | -83.1811
3.4367 | -83.1825
3.437 | -83.1836
3.4373 | -83.1846
3.4377 | -83.1854
3.438 | -83.1865
3.4383 | -83.1866
3.4386 | -83.1859
3.4389 | -83.187
3.4392 | -83.1877
3.4395 | -83.1896
3.4398 | -83.1914
3.4401 | -83.192
3.4404 | -83.1937
3.4408 | -83.1936
3.4411 | -83.194
3.4414 | -83.1985
3.4417 | -83.1989
3.442 | -83.1984
3.4423 | -83.2001
3.4426 | -83.201
3.4429 | -83.2024
3.4432 | -83.2034
3.4436 | -83.2057
3.4439 | -83.2064
3.4442 | -83.2074
3.4445 | -83.2094
3.4448 | -83.2121
3.4451 | -83.214
3.4454 | -83.2156
3.4457 | -83.2177
3.446 | -83.2187
3.4463 | -83.2198
3.4467 | -83.2204
3.447 | -83.2226
3.4473 | -83.2247
3.4476 | -83.2252
3.4479 | -83.2264
3.4482 | -83.2271
3.4485 | -83.2283
3.4488 | -83.2295
3.4491 | -83.2304
3.4495 | -83.2309
3.4498 | -83.2307
3.4501 | -83.2317
3.4504 | -83.2345
3.4507 | -83.2361
3.451 | -83.2383
3.4513 | -83.2404
3.4516 | -83.2418
3.4519 | -83.2431
3.4522 | -83.2432
3.4526 | -83.2418
3.4529 | -83.242
3.4532 | -83.2433
3.4535 | -83.2453
3.4538 | -83.2472
3.4541 | -83.2487
3.4544 | -83.2498
3.4547 | -83.2511
3.455 | -83.2527
3.4554 | -83.2539
3.4557 | -83.2557
3.456 | -83.258
3.4563 | -83.2595
3.4566 | -83.2607
3.4569 | -83.262
3.4572 | -83.2635
3.4575 | -83.2639
3.4578 | -83.2659
3.4581 | -83.2675
3.4585 | -83.2673
3.4588 | -83.2684
3.4591 | -83.2699
3.4594 | -83.2712
3.4597 | -83.2722
3.46 | -83.2732
3.4603 | -83.2754
3.4606 | -83.2776
3.4609 | -83.2782
3.4613 | -83.2803
3.4616 | -83.2823
3.4619 | -83.2839
3.4622 | -83.2858
3.4625 | -83.288
3.4628 | -83.2906
3.4631 | -83.2926
3.4634 | -83.2939
3.4637 | -83.2952
3.464 | -83.2971
3.4644 | -83.2987
3.4647 | -83.2997
3.465 | -83.3007
3.4653 | -83.3019
3.4656 | -83.3038
3.4659 | -83.3051
3.4662 | -83.3066
3.4665 | -83.3078
3.4668 | -83.3092
3.4672 | -83.311
3.4675 | -83.3119
3.4678 | -83.3133
3.4681 | -83.3141
3.4684 | -83.3174
3.4687 | -83.3197
3.469 | -83.3208
3.4693 | -83.3227
3.4696 | -83.3231
3.4699 | -83.3237
3.4703 | -83.3249
3.4706 | -83.3272
3.4709 | -83.3299
3.4712 | -83.3319
3.4715 | -83.3343
3.4718 | -83.3361
3.4721 | -83.3373
3.4724 | -83.3388
3.4727 | -83.3404
3.4731 | -83.3421
3.4734 | -83.3441
3.4737 | -83.346
3.474 | -83.3473
3.4743 | -83.3493
3.4746 | -83.3522
3.4749 | -83.3547
3.4752 | -83.3575
3.4755 | -83.359
3.4758 | -83.3595
3.4762 | -83.3596
3.4765 | -83.3614
3.4768 | -83.363
3.4771 | -83.3645
3.4774 | -83.3661
3.4777 | -83.3678
3.478 | -83.3697
3.4783 | -83.3712
3.4786 | -83.3732
3.479 | -83.3749
3.4793 | -83.377
3.4796 | -83.3788
3.4799 | -83.3814
3.4802 | -83.3845
3.4805 | -83.3861
3.4808 | -83.3881
3.4811 | -83.3902
3.4814 | -83.3913
3.4817 | -83.3933
3.4821 | -83.3954
3.4824 | -83.3971
3.4827 | -83.3975
3.483 | -83.3993
3.4833 | -83.4032
3.4836 | -83.4064
3.4839 | -83.4074
3.4842 | -83.4084
3.4845 | -83.4086
3.4849 | -83.4114
3.4852 | -83.4132
3.4855 | -83.415
3.4858 | -83.4162
3.4861 | -83.419
3.4864 | -83.4216
3.4867 | -83.423
3.487 | -83.4245
3.4873 | -83.4272
3.4876 | -83.4291
3.488 | -83.4332
3.4883 | -83.4352
3.4886 | -83.4352
3.4889 | -83.4355
3.4892 | -83.4366
3.4895 | -83.4379
3.4898 | -83.4391
3.4901 | -83.4403
3.4904 | -83.4418
3.4908 | -83.4434
3.4911 | -83.4453
3.4914 | -83.4488
3.4917 | -83.451
3.492 | -83.4523
3.4923 | -83.4535
3.4926 | -83.4551
3.4929 | -83.4566
3.4932 | -83.4589
3.4935 | -83.4616
3.4939 | -83.4635
3.4942 | -83.4646
3.4945 | -83.4661
3.4948 | -83.4677
3.4951 | -83.4686
3.4954 | -83.4701
3.4957 | -83.4718
3.496 | -83.4711
3.4963 | -83.4711
3.4967 | -83.4729
3.497 | -83.4749
3.4973 | -83.4768
3.4976 | -83.4781
3.4979 | -83.4794
3.4982 | -83.4822
3.4985 | -83.4831
3.4988 | -83.485
3.4991 | -83.4858
3.4994 | -83.4871
3.4998 | -83.4884
3.5001 | -83.4904
3.5004 | -83.4924
3.5007 | -83.4928
3.501 | -83.4924
3.5013 | -83.492
3.5016 | -83.4934
3.5019 | -83.4952
3.5022 | -83.4967
3.5026 | -83.4987
3.5029 | -83.5009
3.5032 | -83.5026
3.5035 | -83.5036
3.5038 | -83.5051
3.5041 | -83.5065
3.5044 | -83.5081
3.5047 | -83.509
3.505 | -83.5102
3.5054 | -83.5115
3.5057 | -83.5126
3.506 | -83.5148
3.5063 | -83.5167
3.5066 | -83.5175
3.5069 | -83.5195
3.5072 | -83.5201
3.5075 | -83.5217
3.5078 | -83.5248
3.5081 | -83.526
3.5085 | -83.5269
3.5088 | -83.5275
3.5091 | -83.5286
3.5094 | -83.5296
3.5097 | -83.5302
3.51 | -83.5319
3.5103 | -83.5331
3.5106 | -83.5347
3.5109 | -83.5373
3.5113 | -83.5401
3.5116 | -83.5434
3.5119 | -83.5466
3.5122 | -83.5481
3.5125 | -83.5491
3.5128 | -83.5501
3.5131 | -83.5516
3.5134 | -83.5533
3.5137 | -83.5543
3.514 | -83.5571
3.5144 | -83.5593
3.5147 | -83.56
3.515 | -83.5598
3.5153 | -83.561
3.5156 | -83.5636
3.5159 | -83.5656
3.5162 | -83.5676
3.5165 | -83.5664
3.5168 | -83.5671
3.5172 | -83.5679
3.5175 | -83.5687
3.5178 | -83.5708
3.5181 | -83.5731
3.5184 | -83.5751
3.5187 | -83.5768
3.519 | -83.5779
3.5193 | -83.5804
3.5196 | -83.5814
3.5199 | -83.5819
3.5203 | -83.5823
3.5206 | -83.5827
3.5209 | -83.5833
3.5212 | -83.5845
3.5215 | -83.5862
3.5218 | -83.5891
3.5221 | -83.5903
3.5224 | -83.5926
3.5227 | -83.5935
3.5231 | -83.5952
3.5234 | -83.5971
3.5237 | -83.5978
3.524 | -83.5987
3.5243 | -83.5997
3.5246 | -83.6008
3.5249 | -83.6023
3.5252 | -83.6032
3.5255 | -83.6044
3.5258 | -83.6054
3.5262 | -83.606
3.5265 | -83.603
3.5268 | -83.6023
3.5271 | -83.6038
3.5274 | -83.6035
3.5277 | -83.6047
3.528 | -83.6078
3.5283 | -83.6087
3.5286 | -83.6108
3.529 | -83.6145
3.5293 | -83.6158
3.5296 | -83.6156
3.5299 | -83.6159
3.5302 | -83.6159
3.5305 | -83.6147
3.5308 | -83.6149
3.5311 | -83.6166
3.5314 | -83.6184
3.5317 | -83.6192
3.5321 | -83.6189
3.5324 | -83.6197
3.5327 | -83.6203
3.533 | -83.6206
3.5333 | -83.6204
3.5336 | -83.6201
3.5339 | -83.6198
3.5342 | -83.6205
3.5345 | -83.6214
3.5349 | -83.6224
3.5352 | -83.6244
3.5355 | -83.6269
3.5358 | -83.6282
3.5361 | -83.6289
3.5364 | -83.6295
3.5367 | -83.6307
3.537 | -83.6312
3.5373 | -83.6318
3.5376 | -83.6323
3.538 | -83.6322
3.5383 | -83.6319
3.5386 | -83.632
3.5389 | -83.6316
3.5392 | -83.6322
3.5395 | -83.6326
3.5398 | -83.6329
3.5401 | -83.6336
3.5404 | -83.6331
3.5408 | -83.6326
3.5411 | -83.6331
3.5414 | -83.633
3.5417 | -83.6336
3.542 | -83.6337
3.5423 | -83.6327
3.5426 | -83.6344
3.5429 | -83.6356
3.5432 | -83.6353
3.5435 | -83.6362
3.5439 | -83.6346
3.5442 | -83.6334
3.5445 | -83.636
3.5448 | -83.6376
3.5451 | -83.6371
3.5454 | -83.6369
3.5457 | -83.637
3.546 | -83.6369
3.5463 | -83.6367
3.5467 | -83.6362
3.547 | -83.6353
3.5473 | -83.6357
3.5476 | -83.6367
3.5479 | -83.6372
3.5482 | -83.6368
3.5485 | -83.6365
3.5488 | -83.6367
3.5491 | -83.6348
3.5494 | -83.6344
3.5498 | -83.6337
3.5501 | -83.6337
3.5504 | -83.6339
3.5507 | -83.6341
3.551 | -83.6345
3.5513 | -83.635
3.5516 | -83.6349
3.5519 | -83.6355
3.5522 | -83.6359
3.5526 | -83.6359
3.5529 | -83.6359
3.5532 | -83.6365
3.5535 | -83.637
3.5538 | -83.6369
3.5541 | -83.6366
3.5544 | -83.6368
3.5547 | -83.6371
3.555 | -83.638
3.5553 | -83.6379
3.5557 | -83.6399
3.556 | -83.6411
3.5563 | -83.6412
3.5566 | -83.641
3.5569 | -83.6409
3.5572 | -83.6401
3.5575 | -83.6401
3.5578 | -83.6403
3.5581 | -83.6404
3.5585 | -83.6411
3.5588 | -83.6419
3.5591 | -83.6414
3.5594 | -83.6412
3.5597 | -83.6425
3.56 | -83.6443
3.5603 | -83.6449
3.5606 | -83.646
3.5609 | -83.6466
3.5612 | -83.6461
3.5616 | -83.6474
3.5619 | -83.6468
3.5622 | -83.6456
3.5625 | -83.6453
3.5628 | -83.6453
3.5631 | -83.6453
3.5634 | -83.6456
3.5637 | -83.6451
3.564 | -83.6439
3.5644 | -83.6441
3.5647 | -83.6445
3.565 | -83.6446
3.5653 | -83.6438
3.5656 | -83.6416
3.5659 | -83.6409
3.5662 | -83.641
3.5665 | -83.6412
3.5668 | -83.6421
3.5671 | -83.6425
3.5675 | -83.6406
3.5678 | -83.64
3.5681 | -83.6398
3.5684 | -83.6404
3.5687 | -83.6392
3.569 | -83.6378
3.5693 | -83.6376
3.5696 | -83.6377
3.5699 | -83.6393
3.5703 | -83.6397
3.5706 | -83.6368
3.5709 | -83.6342
3.5712 | -83.6339
3.5715 | -83.6344
3.5718 | -83.6346
3.5721 | -83.6347
3.5724 | -83.6356
3.5727 | -83.6359
3.573 | -83.6352
3.5734 | -83.6348
3.5737 | -83.6351
3.574 | -83.6356
3.5743 | -83.6348
3.5746 | -83.6343
3.5749 | -83.6346
3.5752 | -83.6352
3.5755 | -83.6355
3.5758 | -83.635
3.5762 | -83.6353
3.5765 | -83.6345
3.5768 | -83.6316
3.5771 | -83.6317
3.5774 | -83.6311
3.5777 | -83.6302
3.578 | -83.6291
3.5783 | -83.6289
3.5786 | -83.6297
3.5789 | -83.6304
3.5793 | -83.6303
3.5796 | -83.6298
3.5799 | -83.6291
3.5802 | -83.6311
3.5805 | -83.6308
3.5808 | -83.6295
3.5811 | -83.6292
3.5814 | -83.628
3.5817 | -83.6281
3.5821 | -83.6296
3.5824 | -83.632
3.5827 | -83.6335
3.583 | -83.6329
3.5833 | -83.6316
3.5836 | -83.6311
3.5839 | -83.6305
3.5842 | -83.6297
3.5845 | -83.6284
3.5848 | -83.6259
3.5852 | -83.625
3.5855 | -83.6249
3.5858 | -83.6231
3.5861 | -83.6212
3.5864 | -83.6206
3.5867 | -83.6207
3.587 | -83.6205
3.5873 | -83.6197
3.5876 | -83.6191
3.588 | -83.6187
3.5883 | -83.618
3.5886 | -83.618
3.5889 | -83.6183
3.5892 | -83.6171
3.5895 | -83.6163
3.5898 | -83.6153
3.5901 | -83.6139
3.5904 | -83.6125
3.5907 | -83.6116
3.5911 | -83.6122
3.5914 | -83.6136
3.5917 | -83.6125
3.592 | -83.6126
3.5923 | -83.6118
3.5926 | -83.6116
3.5929 | -83.6109
3.5932 | -83.6103
3.5935 | -83.6094
3.5939 | -83.608
3.5942 | -83.6063
3.5945 | -83.6059
3.5948 | -83.6054
3.5951 | -83.6041
3.5954 | -83.6041
3.5957 | -83.6043
3.596 | -83.6032
3.5963 | -83.6033
3.5966 | -83.6017
3.597 | -83.6007
3.5973 | -83.601
3.5976 | -83.6003
3.5979 | -83.5997
3.5982 | -83.5988
3.5985 | -83.5965
3.5988 | -83.5963
3.5991 | -83.5958
3.5994 | -83.5946
3.5998 | -83.5931
3.6001 | -83.5911
3.6004 | -83.5901
3.6007 | -83.59
3.601 | -83.5888
3.6013 | -83.587
3.6016 | -83.5862
3.6019 | -83.5842
3.6022 | -83.5825
3.6025 | -83.5812
3.6029 | -83.5785
3.6032 | -83.5764
3.6035 | -83.5769
3.6038 | -83.5768
3.6041 | -83.5751
3.6044 | -83.5737
3.6047 | -83.5736
3.605 | -83.5716
3.6053 | -83.5696
3.6057 | -83.5687
3.606 | -83.5676
3.6063 | -83.5664
3.6066 | -83.5657
3.6069 | -83.5645
3.6072 | -83.5618
3.6075 | -83.5595
3.6078 | -83.5583
3.6081 | -83.557
3.6084 | -83.5563
3.6088 | -83.5554
3.6091 | -83.5545
3.6094 | -83.553
3.6097 | -83.5523
3.61 | -83.5524
3.6103 | -83.5507
3.6106 | -83.5497
3.6109 | -83.5481
3.6112 | -83.547
3.6116 | -83.5457
3.6119 | -83.5449
3.6122 | -83.544
3.6125 | -83.5438
3.6128 | -83.5431
3.6131 | -83.5416
3.6134 | -83.5412
3.6137 | -83.5403
3.614 | -83.5389
3.6143 | -83.5368
3.6147 | -83.5354
3.615 | -83.5346
3.6153 | -83.5331
3.6156 | -83.5315
3.6159 | -83.5303
3.6162 | -83.5291
3.6165 | -83.5288
3.6168 | -83.5288
3.6171 | -83.5289
3.6175 | -83.5285
3.6178 | -83.5277
3.6181 | -83.5265
3.6184 | -83.5262
3.6187 | -83.5252
3.619 | -83.5244
3.6193 | -83.5236
3.6196 | -83.5233
3.6199 | -83.5207
3.6202 | -83.5174
3.6206 | -83.5169
3.6209 | -83.5158
3.6212 | -83.5138
3.6215 | -83.5103
3.6218 | -83.5088
3.6221 | -83.5077
3.6224 | -83.506
3.6227 | -83.5043
3.623 | -83.5034
3.6234 | -83.5017
3.6237 | -83.5001
3.624 | -83.4998
3.6243 | -83.499
3.6246 | -83.4984
3.6249 | -83.4968
3.6252 | -83.4957
3.6255 | -83.4954
3.6258 | -83.4939
3.6261 | -83.4922
3.6265 | -83.4913
3.6268 | -83.4902
3.6271 | -83.4892
3.6274 | -83.4878
3.6277 | -83.4862
3.628 | -83.4842
3.6283 | -83.4822
3.6286 | -83.4811
3.6289 | -83.4801
3.6293 | -83.4789
3.6296 | -83.4786
3.6299 | -83.4781
3.6302 | -83.4765
3.6305 | -83.4754
3.6308 | -83.4743
3.6311 | -83.473
3.6314 | -83.4716
3.6317 | -83.4685
3.6321 | -83.4651
3.6324 | -83.4624
3.6327 | -83.4614
3.633 | -83.4593
3.6333 | -83.4573
3.6336 | -83.4563
3.6339 | -83.4558
3.6342 | -83.4553
3.6345 | -83.4542
3.6348 | -83.4533
3.6352 | -83.452
3.6355 | -83.4519
3.6358 | -83.4524
3.6361 | -83.4519
3.6364 | -83.4506
3.6367 | -83.4491
3.637 | -83.4471
3.6373 | -83.446
3.6376 | -83.446
3.638 | -83.445
3.6383 | -83.4445
3.6386 | -83.4449
3.6389 | -83.4447
3.6392 | -83.4436
3.6395 | -83.4423
3.6398 | -83.4419
3.6401 | -83.4407
3.6404 | -83.4373
3.6407 | -83.4349
3.6411 | -83.4335
3.6414 | -83.4326
3.6417 | -83.431
3.642 | -83.4299
3.6423 | -83.43
3.6426 | -83.4284
3.6429 | -83.4276
3.6432 | -83.4264
3.6435 | -83.4252
3.6439 | -83.4236
3.6442 | -83.4248
3.6445 | -83.4259
3.6448 | -83.4246
3.6451 | -83.423
3.6454 | -83.4217
3.6457 | -83.42
3.646 | -83.4179
3.6463 | -83.4159
3.6466 | -83.4141
3.647 | -83.4123
3.6473 | -83.4103
3.6476 | -83.4097
3.6479 | -83.4112
3.6482 | -83.413
3.6485 | -83.4117
3.6488 | -83.4095
3.6491 | -83.4079
3.6494 | -83.4075
3.6498 | -83.4072
3.6501 | -83.4068
3.6504 | -83.4061
3.6507 | -83.4047
3.651 | -83.4034
3.6513 | -83.4017
3.6516 | -83.4003
3.6519 | -83.3995
3.6522 | -83.3982
3.6525 | -83.3971
3.6529 | -83.3972
3.6532 | -83.3964
3.6535 | -83.3954
3.6538 | -83.3947
3.6541 | -83.3934
3.6544 | -83.3924
3.6547 | -83.3911
3.655 | -83.3898
3.6553 | -83.3892
3.6557 | -83.3877
3.656 | -83.3873
3.6563 | -83.3861
3.6566 | -83.3851
3.6569 | -83.3847
3.6572 | -83.3847
3.6575 | -83.3839
3.6578 | -83.3831
3.6581 | -83.3822
3.6584 | -83.3797
3.6588 | -83.3775
3.6591 | -83.3769
3.6594 | -83.3763
3.6597 | -83.3753
3.66 | -83.3742
3.6603 | -83.3734
3.6606 | -83.3716
3.6609 | -83.3704
3.6612 | -83.3691
3.6616 | -83.3665
3.6619 | -83.3648
3.6622 | -83.3644
3.6625 | -83.364
3.6628 | -83.3648
3.6631 | -83.3646
3.6634 | -83.364
3.6637 | -83.3624
3.664 | -83.3602
3.6643 | -83.359
3.6647 | -83.3572
3.665 | -83.356
3.6653 | -83.3555
3.6656 | -83.3549
3.6659 | -83.3548
3.6662 | -83.3543
3.6665 | -83.3533
3.6668 | -83.3525
3.6671 | -83.3517
3.6675 | -83.3498
3.6678 | -83.3489
3.6681 | -83.3482
3.6684 | -83.3471
3.6687 | -83.3468
3.669 | -83.3465
3.6693 | -83.3463
3.6696 | -83.3467
3.6699 | -83.3463
3.6702 | -83.3459
3.6706 | -83.3449
3.6709 | -83.3444
3.6712 | -83.3442
3.6715 | -83.3433
3.6718 | -83.3427
3.6721 | -83.343
3.6724 | -83.3423
3.6727 | -83.3407
3.673 | -83.3393
3.6734 | -83.3385
3.6737 | -83.3381
3.674 | -83.3376
3.6743 | -83.3373
3.6746 | -83.337
3.6749 | -83.3367
3.6752 | -83.3361
3.6755 | -83.3361
3.6758 | -83.336
3.6761 | -83.3369
3.6765 | -83.3367
3.6768 | -83.3358
3.6771 | -83.336
3.6774 | -83.3367
3.6777 | -83.336
3.678 | -83.3354
3.6783 | -83.3354
3.6786 | -83.3356
3.6789 | -83.3364
3.6793 | -83.3373
3.6796 | -83.3372
3.6799 | -83.336
3.6802 | -83.3347
3.6805 | -83.3346
3.6808 | -83.3346
3.6811 | -83.337
3.6814 | -83.3378
3.6817 | -83.3376
3.682 | -83.3377
3.6824 | -83.3374
3.6827 | -83.3392
3.683 | -83.3392
3.6833 | -83.3377
3.6836 | -83.3378
3.6839 | -83.3379
3.6842 | -83.3377
3.6845 | -83.3383
3.6848 | -83.3383
3.6852 | -83.3383
3.6855 | -83.3379
3.6858 | -83.339
3.6861 | -83.3405
3.6864 | -83.3402
3.6867 | -83.3406
3.687 | -83.3394
3.6873 | -83.3393
3.6876 | -83.3401
3.6879 | -83.34
3.6883 | -83.3412
3.6886 | -83.3419
3.6889 | -83.3428
3.6892 | -83.3429
3.6895 | -83.3433
3.6898 | -83.3432
3.6901 | -83.3436
3.6904 | -83.3435
3.6907 | -83.3427
3.6911 | -83.3423
3.6914 | -83.3428
3.6917 | -83.3425
3.692 | -83.3429
3.6923 | -83.3427
3.6926 | -83.3428
3.6929 | -83.3427
3.6932 | -83.3421
3.6935 | -83.3412
3.6938 | -83.341
3.6942 | -83.3408
3.6945 | -83.3403
3.6948 | -83.3401
3.6951 | -83.3399
3.6954 | -83.3403
3.6957 | -83.3411
3.696 | -83.3411
3.6963 | -83.3416
3.6966 | -83.3418
3.697 | -83.3412
3.6973 | -83.3415
3.6976 | -83.3426
3.6979 | -83.3428
3.6982 | -83.3431
3.6985 | -83.3434
3.6988 | -83.3434
3.6991 | -83.3432
3.6994 | -83.3438
3.6997 | -83.344
3.7001 | -83.3441
3.7004 | -83.3456
3.7007 | -83.346
3.701 | -83.3472
3.7013 | -83.3492
3.7016 | -83.3493
3.7019 | -83.3488
3.7022 | -83.349
3.7025 | -83.3497
3.7029 | -83.3499
3.7032 | -83.3503
3.7035 | -83.3512
3.7038 | -83.3515
3.7041 | -83.3515
3.7044 | -83.352
3.7047 | -83.3523
3.705 | -83.3527
3.7053 | -83.3532
3.7056 | -83.3538
3.706 | -83.3542
3.7063 | -83.3554
3.7066 | -83.3558
3.7069 | -83.3556
3.7072 | -83.3555
3.7075 | -83.3565
3.7078 | -83.3574
3.7081 | -83.358
3.7084 | -83.3583
3.7088 | -83.3591
3.7091 | -83.3601
3.7094 | -83.3611
3.7097 | -83.3629
3.71 | -83.3643
3.7103 | -83.3656
3.7106 | -83.3662
3.7109 | -83.3669
3.7112 | -83.3674
3.7115 | -83.3681
3.7119 | -83.3682
3.7122 | -83.3689
3.7125 | -83.3694
3.7128 | -83.3703
3.7131 | -83.3709
3.7134 | -83.3703
3.7137 | -83.3705
3.714 | -83.3716
3.7143 | -83.3734
3.7147 | -83.3733
3.715 | -83.373
3.7153 | -83.3737
3.7156 | -83.3743
3.7159 | -83.3749
3.7162 | -83.3754
3.7165 | -83.3761
3.7168 | -83.3774
3.7171 | -83.3775
3.7174 | -83.3779
3.7178 | -83.3793
3.7181 | -83.3803
3.7184 | -83.3805
3.7187 | -83.3815
3.719 | -83.3827
3.7193 | -83.3833
3.7196 | -83.3851
3.7199 | -83.3868
3.7202 | -83.3887
3.7206 | -83.3897
3.7209 | -83.3902
3.7212 | -83.3919
3.7215 | -83.3932
3.7218 | -83.3933
3.7221 | -83.3942
3.7224 | -83.3951
3.7227 | -83.3961
3.723 | -83.3966
3.7233 | -83.3968
3.7237 | -83.3979
3.724 | -83.3996
3.7243 | -83.4002
3.7246 | -83.3998
3.7249 | -83.4003
3.7252 | -83.4012
3.7255 | -83.4025
3.7258 | -83.4042
3.7261 | -83.4065
3.7265 | -83.4082
3.7268 | -83.4092
3.7271 | -83.4113
3.7274 | -83.4126
3.7277 | -83.4135
3.728 | -83.4139
3.7283 | -83.4141
3.7286 | -83.4147
3.7289 | -83.4156
3.7292 | -83.4155
3.7296 | -83.4163
3.7299 | -83.4171
3.7302 | -83.4177
3.7305 | -83.4186
3.7308 | -83.419
3.7311 | -83.4202
3.7314 | -83.4221
3.7317 | -83.4222
3.732 | -83.4233
3.7324 | -83.4245
3.7327 | -83.4254
3.733 | -83.4262
3.7333 | -83.4292
3.7336 | -83.4313
3.7339 | -83.4323
3.7342 | -83.4333
3.7345 | -83.4349
3.7348 | -83.4366
3.7351 | -83.4384
3.7355 | -83.4388
3.7358 | -83.4405
3.7361 | -83.4417
3.7364 | -83.4428
3.7367 | -83.4441
3.737 | -83.446
3.7373 | -83.4473
3.7376 | -83.448
3.7379 | -83.4493
3.7383 | -83.4507
3.7386 | -83.4518
3.7389 | -83.4521
3.7392 | -83.4532
3.7395 | -83.4547
3.7398 | -83.4575
3.7401 | -83.4603
3.7404 | -83.4619
3.7407 | -83.463
3.741 | -83.4624
3.7414 | -83.4625
3.7417 | -83.4636
3.742 | -83.4648
3.7423 | -83.4652
3.7426 | -83.4652
3.7429 | -83.4658
3.7432 | -83.4662
3.7435 | -83.4668
3.7438 | -83.4674
3.7442 | -83.4687
3.7445 | -83.4703
3.7448 | -83.4717
3.7451 | -83.4738
3.7454 | -83.4752
3.7457 | -83.4768
3.746 | -83.4784
3.7463 | -83.4794
3.7466 | -83.4808
3.7469 | -83.483
3.7473 | -83.485
3.7476 | -83.4868
3.7479 | -83.4907
3.7482 | -83.4933
3.7485 | -83.4947
3.7488 | -83.496
3.7491 | -83.4969
3.7494 | -83.4983
3.7497 | -83.5003
3.7501 | -83.502
3.7504 | -83.5036
3.7507 | -83.5042
3.751 | -83.5053
3.7513 | -83.5058
3.7516 | -83.5073
3.7519 | -83.5082
3.7522 | -83.5091
3.7525 | -83.5112
3.7528 | -83.5141
3.7532 | -83.5174
3.7535 | -83.5191
3.7538 | -83.5202
3.7541 | -83.5226
3.7544 | -83.5233
3.7547 | -83.5239
3.755 | -83.525
3.7553 | -83.5263
3.7556 | -83.5276
3.756 | -83.5292
3.7563 | -83.5301
3.7566 | -83.531
3.7569 | -83.5325
3.7572 | -83.5346
3.7575 | -83.5368
3.7578 | -83.5388
3.7581 | -83.5409
3.7584 | -83.543
3.7587 | -83.5444
3.7591 | -83.5463
3.7594 | -83.5476
3.7597 | -83.5506
3.76 | -83.5537
3.7603 | -83.5553
3.7606 | -83.5571
3.7609 | -83.5595
3.7612 | -83.562
3.7615 | -83.563
3.7619 | -83.5642
3.7622 | -83.5662
3.7625 | -83.5676
3.7628 | -83.5694
3.7631 | -83.5717
3.7634 | -83.5734
3.7637 | -83.5743
3.764 | -83.5759
3.7643 | -83.5773
3.7647 | -83.5788
3.765 | -83.5807
3.7653 | -83.5822
3.7656 | -83.5849
3.7659 | -83.5866
3.7662 | -83.5868
3.7665 | -83.5885
3.7668 | -83.5884
3.7671 | -83.589
3.7674 | -83.5908
3.7678 | -83.5925
3.7681 | -83.595
3.7684 | -83.5978
3.7687 | -83.5988
3.769 | -83.6011
3.7693 | -83.6031
3.7696 | -83.6064
3.7699 | -83.6075
3.7702 | -83.6093
3.7706 | -83.6112
3.7709 | -83.6125
3.7712 | -83.6137
3.7715 | -83.615
3.7718 | -83.6163
3.7721 | -83.618
3.7724 | -83.6202
3.7727 | -83.6222
3.773 | -83.6251
3.7733 | -83.6275
3.7737 | -83.6294
3.774 | -83.631
3.7743 | -83.6328
3.7746 | -83.6346
3.7749 | -83.6358
3.7752 | -83.6375
3.7755 | -83.6388
3.7758 | -83.6407
3.7761 | -83.6423
3.7765 | -83.6451
3.7768 | -83.6472
3.7771 | -83.6485
3.7774 | -83.6494
3.7777 | -83.651
3.778 | -83.6519
3.7783 | -83.6527
3.7786 | -83.6543
3.7789 | -83.6565
3.7792 | -83.6576
3.7796 | -83.6595
3.7799 | -83.659
3.7802 | -83.6596
3.7805 | -83.6613
3.7808 | -83.6633
3.7811 | -83.6651
3.7814 | -83.668
3.7817 | -83.67
3.782 | -83.6714
3.7824 | -83.6731
3.7827 | -83.6749
3.783 | -83.6769
3.7833 | -83.6796
3.7836 | -83.6813
3.7839 | -83.6828
3.7842 | -83.6844
3.7845 | -83.6867
3.7848 | -83.6901
3.7851 | -83.6922
3.7855 | -83.6937
3.7858 | -83.6954
3.7861 | -83.6965
3.7864 | -83.6985
3.7867 | -83.7003
3.787 | -83.7012
3.7873 | -83.7029
3.7876 | -83.7052
3.7879 | -83.708
3.7883 | -83.7107
3.7886 | -83.713
3.7889 | -83.7146
3.7892 | -83.7162
3.7895 | -83.718
3.7898 | -83.7207
3.7901 | -83.7218
3.7904 | -83.7234
3.7907 | -83.7251
3.791 | -83.7266
3.7914 | -83.7279
3.7917 | -83.7298
3.792 | -83.7323
3.7923 | -83.735
3.7926 | -83.7372
3.7929 | -83.7399
3.7932 | -83.7409
3.7935 | -83.7429
3.7938 | -83.7457
3.7942 | -83.7476
3.7945 | -83.749
3.7948 | -83.751
3.7951 | -83.7526
3.7954 | -83.755
3.7957 | -83.7576
3.796 | -83.7602
3.7963 | -83.7624
3.7966 | -83.7633
3.7969 | -83.7656
3.7973 | -83.7692
3.7976 | -83.7723
3.7979 | -83.7741
3.7982 | -83.7767
3.7985 | -83.7783
3.7988 | -83.7795
3.7991 | -83.7849
3.7994 | -83.7873
3.7997 | -83.7885
3.8001 | -83.7917
3.8004 | -83.7946
3.8007 | -83.7952
3.801 | -83.7948
3.8013 | -83.7953
3.8016 | -83.7968
3.8019 | -83.7989
3.8022 | -83.8016
3.8025 | -83.8034
3.8028 | -83.8036
3.8032 | -83.8041
3.8035 | -83.8061
3.8038 | -83.8081
3.8041 | -83.8086
3.8044 | -83.8093
3.8047 | -83.8122
3.805 | -83.8121
3.8053 | -83.8116
3.8056 | -83.8122
3.806 | -83.8134
3.8063 | -83.8147
3.8066 | -83.8162
3.8069 | -83.8175
3.8072 | -83.8192
3.8075 | -83.8201
3.8078 | -83.8203
3.8081 | -83.8217
3.8084 | -83.8228
3.8087 | -83.8251
3.8091 | -83.8274
3.8094 | -83.8287
3.8097 | -83.8304
3.81 | -83.8323
3.8103 | -83.8339
3.8106 | -83.8357
3.8109 | -83.8371
3.8112 | -83.839
3.8115 | -83.8409
3.8119 | -83.8427
3.8122 | -83.8432
3.8125 | -83.8433
3.8128 | -83.8443
3.8131 | -83.8441
3.8134 | -83.8457
3.8137 | -83.8475
3.814 | -83.8483
3.8143 | -83.8493
3.8146 | -83.8508
3.815 | -83.8523
3.8153 | -83.8548
3.8156 | -83.8569
3.8159 | -83.8585
3.8162 | -83.8599
3.8165 | -83.8617
3.8168 | -83.8637
3.8171 | -83.8646
3.8174 | -83.8663
3.8178 | -83.8681
3.8181 | -83.869
3.8184 | -83.8709
3.8187 | -83.8736
3.819 | -83.8754
3.8193 | -83.8749
3.8196 | -83.8748
3.8199 | -83.8753
3.8202 | -83.8771
3.8205 | -83.878
3.8209 | -83.8769
3.8212 | -83.8785
3.8215 | -83.8795
3.8218 | -83.8819
3.8221 | -83.8852
3.8224 | -83.887
3.8227 | -83.8886
3.823 | -83.8895
3.8233 | -83.8911
3.8237 | -83.8918
3.824 | -83.8934
3.8243 | -83.8969
3.8246 | -83.8991
3.8249 | -83.9002
3.8252 | -83.9015
3.8255 | -83.901
3.8258 | -83.9032
3.8261 | -83.9051
3.8264 | -83.9063
3.8268 | -83.9075
3.8271 | -83.909
3.8274 | -83.911
3.8277 | -83.9154
3.828 | -83.9185
3.8283 | -83.9192
3.8286 | -83.9197
3.8289 | -83.9207
3.8292 | -83.9213
3.8296 | -83.9226
3.8299 | -83.9244
3.8302 | -83.9264
3.8305 | -83.9273
3.8308 | -83.9278
3.8311 | -83.9285
3.8314 | -83.9301
3.8317 | -83.9316
3.832 | -83.9324
3.8323 | -83.9326
3.8327 | -83.9326
3.833 | -83.9342
3.8333 | -83.9357
3.8336 | -83.9358
3.8339 | -83.9374
3.8342 | -83.9415
3.8345 | -83.943
3.8348 | -83.9443
3.8351 | -83.946
3.8355 | -83.9473
3.8358 | -83.9491
3.8361 | -83.951
3.8364 | -83.9521
3.8367 | -83.953
3.837 | -83.9541
3.8373 | -83.9549
3.8376 | -83.9561
3.8379 | -83.9579
3.8382 | -83.9601
3.8386 | -83.962
3.8389 | -83.9632
3.8392 | -83.9653
3.8395 | -83.967
3.8398 | -83.9654
3.8401 | -83.9646
3.8404 | -83.9655
3.8407 | -83.9671
3.841 | -83.9688
3.8414 | -83.9703
3.8417 | -83.9704
3.842 | -83.9714
3.8423 | -83.9724
3.8426 | -83.9741
3.8429 | -83.9769
3.8432 | -83.9776
3.8435 | -83.9762
3.8438 | -83.9773
3.8441 | -83.9794
3.8445 | -83.9815
3.8448 | -83.9825
3.8451 | -83.983
3.8454 | -83.9831
3.8457 | -83.9841
3.846 | -83.985
3.8463 | -83.9851
3.8466 | -83.9857
3.8469 | -83.9857
3.8473 | -83.9866
3.8476 | -83.9871
3.8479 | -83.9884
3.8482 | -83.9898
3.8485 | -83.9908
3.8488 | -83.9921
3.8491 | -83.9926
3.8494 | -83.9932
3.8497 | -83.9929
3.85 | -83.9937
3.8504 | -83.9938
3.8507 | -83.9933
3.851 | -83.994
3.8513 | -83.994
3.8516 | -83.9927
3.8519 | -83.9931
3.8522 | -83.9939
3.8525 | -83.9947
3.8528 | -83.996
3.8532 | -83.9966
3.8535 | -83.9971
3.8538 | -83.998
3.8541 | -83.9992
3.8544 | -84.0006
3.8547 | -84.0019
3.855 | -84.0027
3.8553 | -84.002
3.8556 | -84.0039
3.8559 | -84.0061
3.8563 | -84.0065
3.8566 | -84.0069
3.8569 | -84.0083
3.8572 | -84.0089
3.8575 | -84.0117
3.8578 | -84.0133
3.8581 | -84.0141
3.8584 | -84.0148
3.8587 | -84.0166
3.8591 | -84.0184
3.8594 | -84.0191
3.8597 | -84.0189
3.86 | -84.02
3.8603 | -84.0216
3.8606 | -84.0223
3.8609 | -84.0229
3.8612 | -84.0235
3.8615 | -84.024
3.8618 | -84.0242
3.8622 | -84.0246
3.8625 | -84.0265
3.8628 | -84.0289
3.8631 | -84.0298
3.8634 | -84.0296
3.8637 | -84.0307
3.864 | -84.032
3.8643 | -84.0314
3.8646 | -84.0323
3.865 | -84.0338
3.8653 | -84.0342
3.8656 | -84.0344
3.8659 | -84.0348
3.8662 | -84.0368
3.8665 | -84.0376
3.8668 | -84.0386
3.8671 | -84.04
3.8674 | -84.0393
3.8677 | -84.0375
3.8681 | -84.0372
3.8684 | -84.0371
3.8687 | -84.0363
3.869 | -84.0372
3.8693 | -84.0377
3.8696 | -84.0377
3.8699 | -84.0378
3.8702 | -84.0361
3.8705 | -84.0355
3.8709 | -84.0356
3.8712 | -84.0362
3.8715 | -84.0362
3.8718 | -84.0353
3.8721 | -84.0336
3.8724 | -84.0331
3.8727 | -84.033
3.873 | -84.0333
3.8733 | -84.034
3.8736 | -84.0342
3.874 | -84.0347
3.8743 | -84.0343
3.8746 | -84.0342
3.8749 | -84.0345
3.8752 | -84.0343
3.8755 | -84.0337
3.8758 | -84.0344
3.8761 | -84.0345
3.8764 | -84.0348
3.8768 | -84.0338
3.8771 | -84.0339
3.8774 | -84.0335
3.8777 | -84.0325
3.878 | -84.0327
3.8783 | -84.0323
3.8786 | -84.0325
3.8789 | -84.0331
3.8792 | -84.0335
3.8795 | -84.034
3.8799 | -84.0341
3.8802 | -84.0333
3.8805 | -84.0336
3.8808 | -84.0353
3.8811 | -84.0372
3.8814 | -84.0372
3.8817 | -84.037
3.882 | -84.0361
3.8823 | -84.0354
3.8827 | -84.031
3.883 | -84.0286
3.8833 | -84.0276
3.8836 | -84.0283
3.8839 | -84.0276
3.8842 | -84.0271
3.8845 | -84.0284
3.8848 | -84.0291
3.8851 | -84.0281
3.8854 | -84.028
3.8858 | -84.0285
3.8861 | -84.029
3.8864 | -84.0289
3.8867 | -84.029
3.887 | -84.0293
3.8873 | -84.029
3.8876 | -84.0288
3.8879 | -84.0287
3.8882 | -84.0284
3.8886 | -84.0276
3.8889 | -84.0268
3.8892 | -84.0271
3.8895 | -84.0274
3.8898 | -84.0271
3.8901 | -84.0268
3.8904 | -84.0261
3.8907 | -84.026
3.891 | -84.0254
3.8913 | -84.0252
3.8917 | -84.0254
3.892 | -84.0259
3.8923 | -84.0262
3.8926 | -84.0266
3.8929 | -84.026
3.8932 | -84.0261
3.8935 | -84.0261
3.8938 | -84.028
3.8941 | -84.0268
3.8945 | -84.0276
3.8948 | -84.0271
3.8951 | -84.0251
3.8954 | -84.0244
3.8957 | -84.024
3.896 | -84.0239
3.8963 | -84.0234
3.8966 | -84.0228
3.8969 | -84.0225
3.8972 | -84.0223
3.8976 | -84.0213
3.8979 | -84.0199
3.8982 | -84.0196
3.8985 | -84.0195
3.8988 | -84.0201
3.8991 | -84.0198
3.8994 | -84.0188
3.8997 | -84.0182
3.9 | -84.0177
3.9004 | -84.0172
3.9007 | -84.0169
3.901 | -84.0162
3.9013 | -84.0157
3.9016 | -84.0151
3.9019 | -84.0141
3.9022 | -84.0132
3.9025 | -84.0153
3.9028 | -84.0159
3.9031 | -84.0149
3.9035 | -84.0148
3.9038 | -84.0152
3.9041 | -84.0146
3.9044 | -84.0138
3.9047 | -84.0137
3.905 | -84.0134
3.9053 | -84.0131
3.9056 | -84.0127
3.9059 | -84.0127
3.9063 | -84.0128
3.9066 | -84.013
3.9069 | -84.0117
3.9072 | -84.0091
3.9075 | -84.0084
3.9078 | -84.0081
3.9081 | -84.0088
3.9084 | -84.0104
3.9087 | -84.0112
3.9091 | -84.011
3.9094 | -84.0117
3.9097 | -84.0119
3.91 | -84.0119
3.9103 | -84.0117
3.9106 | -84.0127
3.9109 | -84.0128
3.9112 | -84.0118
3.9115 | -84.0113
3.9118 | -84.0105
3.9122 | -84.0098
3.9125 | -84.0103
3.9128 | -84.0104
3.9131 | -84.0097
3.9134 | -84.0085
3.9137 | -84.0071
3.914 | -84.0053
3.9143 | -84.0043
3.9146 | -84.0027
3.915 | -84.0008
3.9153 | -84.0012
3.9156 | -84.0006
3.9159 | -83.9995
3.9162 | -83.9985
3.9165 | -83.9979
3.9168 | -83.9976
3.9171 | -83.997
3.9174 | -83.995
3.9177 | -83.9949
3.9181 | -83.9942
3.9184 | -83.9931
3.9187 | -83.9934
3.919 | -83.9927
3.9193 | -83.9916
3.9196 | -83.9914
3.9199 | -83.9905
3.9202 | -83.9896
3.9205 | -83.9889
3.9209 | -83.9881
3.9212 | -83.9876
3.9215 | -83.9872
3.9218 | -83.9885
3.9221 | -83.9893
3.9224 | -83.9894
3.9227 | -83.9892
3.923 | -83.9873
3.9233 | -83.9849
3.9236 | -83.9841
3.924 | -83.9846
3.9243 | -83.9844
3.9246 | -83.984
3.9249 | -83.984
3.9252 | -83.985
3.9255 | -83.9849
3.9258 | -83.984
3.9261 | -83.9829
3.9264 | -83.982
3.9268 | -83.9797
3.9271 | -83.9763
3.9274 | -83.9759
3.9277 | -83.9765
3.928 | -83.9771
3.9283 | -83.9777
3.9286 | -83.9784
3.9289 | -83.9785
3.9292 | -83.9758
3.9295 | -83.9727
3.9299 | -83.9712
3.9302 | -83.9706
3.9305 | -83.9704
3.9308 | -83.9701
3.9311 | -83.9697
3.9314 | -83.9707
3.9317 | -83.971
3.932 | -83.9694
3.9323 | -83.9696
3.9327 | -83.9688
3.933 | -83.9666
3.9333 | -83.9651
3.9336 | -83.9643
3.9339 | -83.9631
3.9342 | -83.963
3.9345 | -83.9618
3.9348 | -83.9614
3.9351 | -83.9601
3.9354 | -83.9591
3.9358 | -83.9586
3.9361 | -83.9573
3.9364 | -83.9569
3.9367 | -83.9563
3.937 | -83.9569
3.9373 | -83.9565
3.9376 | -83.9564
3.9379 | -83.9564
3.9382 | -83.9559
3.9386 | -83.9554
3.9389 | -83.955
3.9392 | -83.9545
3.9395 | -83.9543
3.9398 | -83.9543
3.9401 | -83.9536
3.9404 | -83.953
3.9407 | -83.9524
3.941 | -83.952
3.9413 | -83.9519
3.9417 | -83.9514
3.942 | -83.9511
3.9423 | -83.9517
3.9426 | -83.9514
3.9429 | -83.9515
3.9432 | -83.951
3.9435 | -83.9497
3.9438 | -83.9486
3.9441 | -83.948
3.9445 | -83.9478
3.9448 | -83.9487
3.9451 | -83.9491
3.9454 | -83.9492
3.9457 | -83.9489
3.946 | -83.9485
3.9463 | -83.9479
3.9466 | -83.9458
3.9469 | -83.945
3.9472 | -83.9452
3.9476 | -83.9445
3.9479 | -83.9439
3.9482 | -83.9445
3.9485 | -83.945
3.9488 | -83.9423
3.9491 | -83.941
3.9494 | -83.9401
3.9497 | -83.939
3.95 | -83.9382
3.9504 | -83.9377
3.9507 | -83.9375
3.951 | -83.9381
3.9513 | -83.9379
3.9516 | -83.9375
3.9519 | -83.9373
3.9522 | -83.9369
3.9525 | -83.938
3.9528 | -83.9375
3.9531 | -83.936
3.9535 | -83.9342
3.9538 | -83.9322
3.9541 | -83.9313
3.9544 | -83.931
3.9547 | -83.9305
3.955 | -83.9312
3.9553 | -83.9327
3.9556 | -83.9328
3.9559 | -83.9327
3.9563 | -83.933
3.9566 | -83.9336
3.9569 | -83.9333
3.9572 | -83.9318
3.9575 | -83.9317
3.9578 | -83.9314
3.9581 | -83.9307
3.9584 | -83.93
3.9587 | -83.9286
3.959 | -83.9286
3.9594 | -83.9293
3.9597 | -83.9291
3.96 | -83.9279
3.9603 | -83.9272
3.9606 | -83.9269
3.9609 | -83.9275
3.9612 | -83.9272
3.9615 | -83.9262
3.9618 | -83.9258
3.9622 | -83.9257
3.9625 | -83.9251
3.9628 | -83.9244
3.9631 | -83.9245
3.9634 | -83.9244
3.9637 | -83.9234
3.964 | -83.9235
3.9643 | -83.9246
3.9646 | -83.9254
3.9649 | -83.9259
3.9653 | -83.9257
3.9656 | -83.9253
3.9659 | -83.9249
3.9662 | -83.9251
3.9665 | -83.9249
3.9668 | -83.9248
3.9671 | -83.9244
3.9674 | -83.9243
3.9677 | -83.9249
3.9681 | -83.9256
3.9684 | -83.9248
3.9687 | -83.924
3.969 | -83.9221
3.9693 | -83.9175
3.9696 | -83.9164
3.9699 | -83.9162
3.9702 | -83.9154
3.9705 | -83.915
3.9708 | -83.9151
3.9712 | -83.9137
3.9715 | -83.9123
3.9718 | -83.9121
3.9721 | -83.912
3.9724 | -83.9118
3.9727 | -83.9121
3.973 | -83.9121
3.9733 | -83.9116
3.9736 | -83.9122
3.974 | -83.9135
3.9743 | -83.9143
3.9746 | -83.9139
3.9749 | -83.9138
3.9752 | -83.9139
3.9755 | -83.914
3.9758 | -83.9148
3.9761 | -83.9143
3.9764 | -83.9128
3.9767 | -83.911
3.9771 | -83.9097
3.9774 | -83.9088
3.9777 | -83.9084
3.978 | -83.9084
3.9783 | -83.9079
3.9786 | -83.9076
3.9789 | -83.9075
3.9792 | -83.9077
3.9795 | -83.9071
3.9799 | -83.9062
3.9802 | -83.906
3.9805 | -83.9058
3.9808 | -83.9069
3.9811 | -83.9079
3.9814 | -83.9086
3.9817 | -83.9088
3.982 | -83.9101
3.9823 | -83.9098
3.9826 | -83.9097
3.983 | -83.9099
3.9833 | -83.9108
3.9836 | -83.9112
3.9839 | -83.9112
3.9842 | -83.9116
3.9845 | -83.9127
3.9848 | -83.913
3.9851 | -83.913
3.9854 | -83.9136
3.9858 | -83.9135
3.9861 | -83.9133
3.9864 | -83.9136
3.9867 | -83.9137
3.987 | -83.9149
3.9873 | -83.9144
3.9876 | -83.9139
3.9879 | -83.9134
3.9882 | -83.9135
3.9885 | -83.9134
3.9889 | -83.9131
3.9892 | -83.9129
3.9895 | -83.9127
3.9898 | -83.9133
3.9901 | -83.9136
3.9904 | -83.9131
3.9907 | -83.9125
3.991 | -83.9123
3.9913 | -83.9129
3.9917 | -83.9126
3.992 | -83.9127
3.9923 | -83.9127
3.9926 | -83.9129
3.9929 | -83.9127
3.9932 | -83.9128
3.9935 | -83.9129
3.9938 | -83.9134
3.9941 | -83.9138
3.9944 | -83.9138
3.9948 | -83.9137
3.9951 | -83.9136
3.9954 | -83.9139
3.9957 | -83.9139
3.996 | -83.9148
3.9963 | -83.9152
3.9966 | -83.9161
3.9969 | -83.917
3.9972 | -83.9167
3.9976 | -83.9167
3.9979 | -83.9177
3.9982 | -83.9179
3.9985 | -83.9171
3.9988 | -83.9163
3.9991 | -83.9164
3.9994 | -83.917
3.9997 | -83.9167
4 | -83.9168
4.0003 | -83.9186
4.0007 | -83.919
4.001 | -83.9182
4.0013 | -83.9181
4.0016 | -83.9189
4.0019 | -83.9186
4.0022 | -83.9181
4.0025 | -83.9185
4.0028 | -83.9189
4.0031 | -83.919
4.0035 | -83.9185
4.0038 | -83.9189
4.0041 | -83.9195
4.0044 | -83.9195
4.0047 | -83.92
4.005 | -83.9205
4.0053 | -83.921
4.0056 | -83.9207
4.0059 | -83.9189
4.0062 | -83.9175
4.0066 | -83.9168
4.0069 | -83.9164
4.0072 | -83.9163
4.0075 | -83.9159
4.0078 | -83.915
4.0081 | -83.9157
4.0084 | -83.9157
4.0087 | -83.916
4.009 | -83.9159
4.0094 | -83.9165
4.0097 | -83.9166
4.01 | -83.9172
4.0103 | -83.9195
4.0106 | -83.9211
4.0109 | -83.9215
4.0112 | -83.9212
4.0115 | -83.9213
4.0118 | -83.9228
4.0121 | -83.9236
4.0125 | -83.9245
4.0128 | -83.9253
4.0131 | -83.9259
4.0134 | -83.9261
4.0137 | -83.9273
4.014 | -83.9267
4.0143 | -83.9253
4.0146 | -83.9248
4.0149 | -83.9252
4.0153 | -83.9258
4.0156 | -83.9264
4.0159 | -83.927
4.0162 | -83.9272
4.0165 | -83.9269
4.0168 | -83.9271
4.0171 | -83.9272
4.0174 | -83.9276
4.0177 | -83.9277
4.018 | -83.928
4.0184 | -83.9284
4.0187 | -83.9282
4.019 | -83.9289
4.0193 | -83.9297
4.0196 | -83.9296
4.0199 | -83.9299
4.0202 | -83.9308
4.0205 | -83.9319
4.0208 | -83.9334
4.0212 | -83.935
4.0215 | -83.9356
4.0218 | -83.9368
4.0221 | -83.9383
4.0224 | -83.9393
4.0227 | -83.9392
4.023 | -83.9391
4.0233 | -83.9396
4.0236 | -83.9392
4.0239 | -83.9398
4.0243 | -83.9403
4.0246 | -83.9406
4.0249 | -83.9416
4.0252 | -83.943
4.0255 | -83.944
4.0258 | -83.9447
4.0261 | -83.945
4.0264 | -83.9456
4.0267 | -83.9465
4.0271 | -83.9479
4.0274 | -83.9481
4.0277 | -83.9478
4.028 | -83.9483
4.0283 | -83.9491
4.0286 | -83.9497
4.0289 | -83.9493
4.0292 | -83.9477
4.0295 | -83.9481
4.0298 | -83.9484
4.0302 | -83.9478
4.0305 | -83.948
4.0308 | -83.9485
4.0311 | -83.9489
4.0314 | -83.9497
4.0317 | -83.9513
4.032 | -83.9516
4.0323 | -83.9531
4.0326 | -83.9538
4.033 | -83.9546
4.0333 | -83.9553
4.0336 | -83.956
4.0339 | -83.957
4.0342 | -83.9576
4.0345 | -83.9589
4.0348 | -83.9604
4.0351 | -83.961
4.0354 | -83.962
4.0357 | -83.9625
4.0361 | -83.9628
4.0364 | -83.9632
4.0367 | -83.9635
4.037 | -83.965
4.0373 | -83.9671
4.0376 | -83.9672
4.0379 | -83.9694
4.0382 | -83.9711
4.0385 | -83.9714
4.0389 | -83.9724
4.0392 | -83.9735
4.0395 | -83.9741
4.0398 | -83.9757
4.0401 | -83.9763
4.0404 | -83.9776
4.0407 | -83.979
4.041 | -83.9792
4.0413 | -83.9793
4.0416 | -83.98
4.042 | -83.9807
4.0423 | -83.9801
4.0426 | -83.9792
4.0429 | -83.9806
4.0432 | -83.9819
4.0435 | -83.9833
4.0438 | -83.9841
4.0441 | -83.9849
4.0444 | -83.9858
4.0448 | -83.9867
4.0451 | -83.9873
4.0454 | -83.9883
4.0457 | -83.9897
4.046 | -83.991
4.0463 | -83.992
4.0466 | -83.9934
4.0469 | -83.994
4.0472 | -83.9943
4.0476 | -83.9954
4.0479 | -83.9964
4.0482 | -83.9974
4.0485 | -83.9983
4.0488 | -83.9976
4.0491 | -83.9978
4.0494 | -83.9977
4.0497 | -83.9996
4.05 | -83.9997
4.0503 | -84.0004
4.0507 | -84.0018
4.051 | -84.0016
4.0513 | -84.0019
4.0516 | -84.0034
4.0519 | -84.0054
4.0522 | -84.0091
4.0525 | -84.0103
4.0528 | -84.0108
4.0531 | -84.0115
4.0534 | -84.0121
4.0538 | -84.0129
4.0541 | -84.014
4.0544 | -84.0148
4.0547 | -84.0156
4.055 | -84.0164
4.0553 | -84.0168
4.0556 | -84.018
4.0559 | -84.0187
4.0562 | -84.02
4.0566 | -84.0212
4.0569 | -84.022
4.0572 | -84.0229
4.0575 | -84.0228
4.0578 | -84.0224
4.0581 | -84.0235
4.0584 | -84.0257
4.0587 | -84.0269
4.059 | -84.0274
4.0594 | -84.0281
4.0597 | -84.0297
4.06 | -84.0298
4.0603 | -84.0297
4.0606 | -84.0305
4.0609 | -84.0313
4.0612 | -84.0316
4.0615 | -84.0306
4.0618 | -84.0304
4.0621 | -84.0314
4.0625 | -84.032
4.0628 | -84.0327
4.0631 | -84.034
4.0634 | -84.0348
4.0637 | -84.0347
4.064 | -84.0359
4.0643 | -84.037
4.0646 | -84.038
4.0649 | -84.0384
4.0652 | -84.0381
4.0656 | -84.0384
4.0659 | -84.039
4.0662 | -84.0413
4.0665 | -84.0422
4.0668 | -84.0418
4.0671 | -84.0427
4.0674 | -84.0435
4.0677 | -84.0451
4.068 | -84.0461
4.0684 | -84.0471
4.0687 | -84.0491
4.069 | -84.0499
4.0693 | -84.0509
4.0696 | -84.0518
4.0699 | -84.0539
4.0702 | -84.0545
4.0705 | -84.0556
4.0708 | -84.0564
4.0712 | -84.0572
4.0715 | -84.0582
4.0718 | -84.0608
4.0721 | -84.0623
4.0724 | -84.064
4.0727 | -84.0654
4.073 | -84.0661
4.0733 | -84.0652
4.0736 | -84.065
4.0739 | -84.0665
4.0743 | -84.0673
4.0746 | -84.0693
4.0749 | -84.0713
4.0752 | -84.0723
4.0755 | -84.0733
4.0758 | -84.0742
4.0761 | -84.0748
4.0764 | -84.0725
4.0767 | -84.0748
4.0771 | -84.0768
4.0774 | -84.0773
4.0777 | -84.0774
4.078 | -84.0792
4.0783 | -84.0807
4.0786 | -84.0825
4.0789 | -84.0864
4.0792 | -84.0901
4.0795 | -84.0916
4.0798 | -84.0914
4.0802 | -84.0922
4.0805 | -84.0931
4.0808 | -84.0938
4.0811 | -84.0946
4.0814 | -84.0961
4.0817 | -84.0984
4.082 | -84.0998
4.0823 | -84.1006
4.0826 | -84.1018
4.083 | -84.1026
4.0833 | -84.1032
4.0836 | -84.1047
4.0839 | -84.106
4.0842 | -84.1068
4.0845 | -84.1078
4.0848 | -84.108
4.0851 | -84.1086
4.0854 | -84.1071
4.0857 | -84.1075
4.0861 | -84.1084
4.0864 | -84.1094
4.0867 | -84.1113
4.087 | -84.113
4.0873 | -84.1137
4.0876 | -84.1151
4.0879 | -84.1168
4.0882 | -84.1167
4.0885 | -84.1173
4.0889 | -84.1179
4.0892 | -84.1179
4.0895 | -84.1193
4.0898 | -84.1194
4.0901 | -84.1202
4.0904 | -84.121
4.0907 | -84.1217
4.091 | -84.1222
4.0913 | -84.1231
4.0916 | -84.1242
4.092 | -84.1256
4.0923 | -84.1268
4.0926 | -84.1274
4.0929 | -84.1284
4.0932 | -84.1296
4.0935 | -84.1304
4.0938 | -84.1302
4.0941 | -84.1302
4.0944 | -84.1311
4.0948 | -84.1334
4.0951 | -84.135
4.0954 | -84.136
4.0957 | -84.1385
4.096 | -84.142
4.0963 | -84.1418
4.0966 | -84.1407
4.0969 | -84.1397
4.0972 | -84.1407
4.0975 | -84.142
4.0979 | -84.143
4.0982 | -84.1442
4.0985 | -84.145
4.0988 | -84.1469
4.0991 | -84.1481
4.0994 | -84.1487
4.0997 | -84.1485
4.1 | -84.1496
4.1003 | -84.1489
4.1007 | -84.1477
4.101 | -84.1494
4.1013 | -84.1498
4.1016 | -84.1508
4.1019 | -84.1518
4.1022 | -84.1525
4.1025 | -84.1535
4.1028 | -84.1544
4.1031 | -84.155
4.1034 | -84.156
4.1038 | -84.1574
4.1041 | -84.1592
4.1044 | -84.1608
4.1047 | -84.1615
4.105 | -84.1611
4.1053 | -84.1608
4.1056 | -84.1614
4.1059 | -84.1621
4.1062 | -84.163
4.1066 | -84.1643
4.1069 | -84.1648
4.1072 | -84.164
4.1075 | -84.1644
4.1078 | -84.1647
4.1081 | -84.1655
4.1084 | -84.1672
4.1087 | -84.1678
4.109 | -84.1685
4.1093 | -84.1696
4.1097 | -84.1705
4.11 | -84.1707
4.1103 | -84.1709
4.1106 | -84.1714
4.1109 | -84.1721
4.1112 | -84.1726
4.1115 | -84.1736
4.1118 | -84.1745
4.1121 | -84.1757
4.1125 | -84.1757
4.1128 | -84.1751
4.1131 | -84.1755
4.1134 | -84.1762
4.1137 | -84.1766
4.114 | -84.1772
4.1143 | -84.1778
4.1146 | -84.1782
4.1149 | -84.1783
4.1152 | -84.179
4.1156 | -84.1793
4.1159 | -84.1801
4.1162 | -84.1837
4.1165 | -84.186
4.1168 | -84.1851
4.1171 | -84.1846
4.1174 | -84.1852
4.1177 | -84.186
4.118 | -84.1872
4.1184 | -84.1871
4.1187 | -84.1873
4.119 | -84.1877
4.1193 | -84.1873
4.1196 | -84.1881
4.1199 | -84.1904
4.1202 | -84.1912
4.1205 | -84.1912
4.1208 | -84.192
4.1211 | -84.1924
4.1215 | -84.193
4.1218 | -84.1934
4.1221 | -84.1939
4.1224 | -84.1954
4.1227 | -84.1961
4.123 | -84.1965
4.1233 | -84.1977
4.1236 | -84.1984
4.1239 | -84.1997
4.1243 | -84.1998
4.1246 | -84.2011
4.1249 | -84.201
4.1252 | -84.2009
4.1255 | -84.2019
4.1258 | -84.2022
4.1261 | -84.2032
4.1264 | -84.2036
4.1267 | -84.2045
4.127 | -84.2052
4.1274 | -84.2065
4.1277 | -84.2085
4.128 | -84.2093
4.1283 | -84.2106
4.1286 | -84.2128
4.1289 | -84.2139
4.1292 | -84.2144
4.1295 | -84.2153
4.1298 | -84.2158
4.1302 | -84.2144
4.1305 | -84.2149
4.1308 | -84.2169
4.1311 | -84.217
4.1314 | -84.2176
4.1317 | -84.2173
4.132 | -84.2162
4.1323 | -84.2169
4.1326 | -84.2176
4.1329 | -84.2189
4.1333 | -84.2202
4.1336 | -84.2205
4.1339 | -84.2208
4.1342 | -84.2208
4.1345 | -84.2211
4.1348 | -84.2212
4.1351 | -84.2214
4.1354 | -84.2218
4.1357 | -84.2221
4.1361 | -84.2226
4.1364 | -84.2227
4.1367 | -84.2234
4.137 | -84.2239
4.1373 | -84.2236
4.1376 | -84.2228
4.1379 | -84.2226
4.1382 | -84.2211
4.1385 | -84.2199
4.1388 | -84.2212
4.1392 | -84.2178
4.1395 | -84.2179
4.1398 | -84.2184
4.1401 | -84.219
4.1404 | -84.2196
4.1407 | -84.2196
4.141 | -84.219
4.1413 | -84.2196
4.1416 | -84.2207
4.142 | -84.2217
4.1423 | -84.2215
4.1426 | -84.2224
4.1429 | -84.223
4.1432 | -84.2229
4.1435 | -84.2233
4.1438 | -84.223
4.1441 | -84.2226
4.1444 | -84.2221
4.1447 | -84.2228
4.1451 | -84.2237
4.1454 | -84.2241
4.1457 | -84.2231
4.146 | -84.2229
4.1463 | -84.2229
4.1466 | -84.2234
4.1469 | -84.2249
4.1472 | -84.2249
4.1475 | -84.2245
4.1479 | -84.2247
4.1482 | -84.2247
4.1485 | -84.2256
4.1488 | -84.2275
4.1491 | -84.2284
4.1494 | -84.2288
4.1497 | -84.2289
4.15 | -84.2291
4.1503 | -84.2292
4.1506 | -84.2305
4.151 | -84.2321
4.1513 | -84.2326
4.1516 | -84.2324
4.1519 | -84.2322
4.1522 | -84.2314
4.1525 | -84.2314
4.1528 | -84.232
4.1531 | -84.233
4.1534 | -84.2331
4.1538 | -84.2331
4.1541 | -84.2332
4.1544 | -84.2337
4.1547 | -84.2333
4.155 | -84.233
4.1553 | -84.2328
4.1556 | -84.2343
4.1559 | -84.2344
4.1562 | -84.2352
4.1565 | -84.2361
4.1569 | -84.2361
4.1572 | -84.2368
4.1575 | -84.2374
4.1578 | -84.2375
4.1581 | -84.2374
4.1584 | -84.2373
4.1587 | -84.2373
4.159 | -84.2366
4.1593 | -84.2355
4.1597 | -84.2362
4.16 | -84.2359
4.1603 | -84.2365
4.1606 | -84.2368
4.1609 | -84.2356
4.1612 | -84.2356
4.1615 | -84.2374
4.1618 | -84.2407
4.1621 | -84.2453
4.1624 | -84.2473
4.1628 | -84.2488
4.1631 | -84.2489
4.1634 | -84.2493
4.1637 | -84.2503
4.164 | -84.25
4.1643 | -84.2486
4.1646 | -84.2492
4.1649 | -84.2493
4.1652 | -84.2489
4.1656 | -84.2461
4.1659 | -84.2419
4.1662 | -84.2402
4.1665 | -84.2397
4.1668 | -84.2396
4.1671 | -84.2399
4.1674 | -84.2397
4.1677 | -84.2392
4.168 | -84.2386
4.1683 | -84.2389
4.1687 | -84.2396
4.169 | -84.2408
4.1693 | -84.2406
4.1696 | -84.2401
4.1699 | -84.2387
4.1702 | -84.2384
4.1705 | -84.2384
4.1708 | -84.2386
4.1711 | -84.2394
4.1715 | -84.239
4.1718 | -84.239
4.1721 | -84.2389
4.1724 | -84.2391
4.1727 | -84.239
4.173 | -84.2385
4.1733 | -84.2396
4.1736 | -84.2419
4.1739 | -84.2385
4.1742 | -84.2363
4.1746 | -84.2362
4.1749 | -84.2373
4.1752 | -84.2377
4.1755 | -84.2389
4.1758 | -84.2385
4.1761 | -84.2383
4.1764 | -84.2369
4.1767 | -84.2365
4.177 | -84.2363
4.1774 | -84.2362
4.1777 | -84.2364
4.178 | -84.2363
4.1783 | -84.2367
4.1786 | -84.2371
4.1789 | -84.2356
4.1792 | -84.2344
4.1795 | -84.2325
4.1798 | -84.2318
4.1801 | -84.2315
4.1805 | -84.2305
4.1808 | -84.2302
4.1811 | -84.2296
4.1814 | -84.229
4.1817 | -84.2287
4.182 | -84.2283
4.1823 | -84.2287
4.1826 | -84.2294
4.1829 | -84.2296
4.1833 | -84.229
4.1836 | -84.2285
4.1839 | -84.2287
4.1842 | -84.2285
4.1845 | -84.2276
4.1848 | -84.2272
4.1851 | -84.2271
4.1854 | -84.2263
4.1857 | -84.2264
4.186 | -84.2246
4.1864 | -84.2228
4.1867 | -84.2228
4.187 | -84.2231
4.1873 | -84.2232
4.1876 | -84.2228
4.1879 | -84.2224
4.1882 | -84.2237
4.1885 | -84.2264
4.1888 | -84.2273
4.1892 | -84.227
4.1895 | -84.2247
4.1898 | -84.2236
4.1901 | -84.2238
4.1904 | -84.2234
4.1907 | -84.2234
4.191 | -84.2229
4.1913 | -84.2218
4.1916 | -84.2212
4.192 | -84.2207
4.1923 | -84.221
4.1926 | -84.2216
4.1929 | -84.2221
4.1932 | -84.2211
4.1935 | -84.2204
4.1938 | -84.2201
4.1941 | -84.2198
4.1944 | -84.2193
4.1947 | -84.2189
4.1951 | -84.2196
4.1954 | -84.22
4.1957 | -84.221
4.196 | -84.2215
4.1963 | -84.2222
4.1966 | -84.2228
4.1969 | -84.2221
4.1972 | -84.2213
4.1975 | -84.22
4.1978 | -84.2196
4.1982 | -84.2194
4.1985 | -84.2168
4.1988 | -84.2158
4.1991 | -84.216
4.1994 | -84.2172
4.1997 | -84.2169
4.2 | -84.2168
4.2003 | -84.2173
4.2006 | -84.217
4.201 | -84.2176
4.2013 | -84.2178
4.2016 | -84.2176
4.2019 | -84.2164
4.2022 | -84.2161
4.2025 | -84.2147
4.2028 | -84.2143
4.2031 | -84.2143
4.2034 | -84.2139
4.2038 | -84.2139
4.2041 | -84.2137
4.2044 | -84.2136
4.2047 | -84.2136
4.205 | -84.2134
4.2053 | -84.2145
4.2056 | -84.2162
4.2059 | -84.2161
4.2062 | -84.2158
4.2065 | -84.2162
4.2069 | -84.2163
4.2072 | -84.2159
4.2075 | -84.2152
4.2078 | -84.2155
4.2081 | -84.2159
4.2084 | -84.2159
4.2087 | -84.216
4.209 | -84.2163
4.2093 | -84.2174
4.2096 | -84.2193
4.21 | -84.2197
4.2103 | -84.22
4.2106 | -84.2191
4.2109 | -84.2196
4.2112 | -84.2186
4.2115 | -84.2178
4.2118 | -84.2176
4.2121 | -84.2195
4.2124 | -84.2214
4.2128 | -84.222
4.2131 | -84.2218
4.2134 | -84.2211
4.2137 | -84.2201
4.214 | -84.2203
4.2143 | -84.22
4.2146 | -84.2201
4.2149 | -84.2214
4.2152 | -84.2215
4.2156 | -84.2211
4.2159 | -84.2217
4.2162 | -84.2222
4.2165 | -84.2216
4.2168 | -84.2215
4.2171 | -84.2218
4.2174 | -84.2216
4.2177 | -84.2222
4.218 | -84.2221
4.2183 | -84.2228
4.2187 | -84.2234
4.219 | -84.2236
4.2193 | -84.2242
4.2196 | -84.2241
4.2199 | -84.2249
4.2202 | -84.2256
4.2205 | -84.225
4.2208 | -84.2255
4.2211 | -84.2262
4.2215 | -84.226
4.2218 | -84.226
4.2221 | -84.2259
4.2224 | -84.2254
4.2227 | -84.2248
4.223 | -84.2248
4.2233 | -84.2244
4.2236 | -84.2253
4.2239 | -84.2255
4.2242 | -84.2254
4.2246 | -84.2251
4.2249 | -84.2246
4.2252 | -84.2248
4.2255 | -84.2247
4.2258 | -84.2248
4.2261 | -84.2249
4.2264 | -84.2247
4.2267 | -84.2247
4.227 | -84.2244
4.2274 | -84.2249
4.2277 | -84.2259
4.228 | -84.2266
4.2283 | -84.2269
4.2286 | -84.227
4.2289 | -84.2279
4.2292 | -84.2283
4.2295 | -84.2278
4.2298 | -84.227
4.2301 | -84.2272
4.2305 | -84.2275
4.2308 | -84.2294
4.2311 | -84.2313
4.2314 | -84.2325
4.2317 | -84.2331
4.232 | -84.2333
4.2323 | -84.2333
4.2326 | -84.2337
4.2329 | -84.236
4.2333 | -84.2369
4.2336 | -84.2375
4.2339 | -84.238
4.2342 | -84.2378
4.2345 | -84.2377
4.2348 | -84.2365
4.2351 | -84.2355
4.2354 | -84.2357
4.2357 | -84.2357
4.236 | -84.2353
4.2364 | -84.2352
4.2367 | -84.2354
4.237 | -84.2359
4.2373 | -84.2366
4.2376 | -84.2374
4.2379 | -84.2357
4.2382 | -84.2371
4.2385 | -84.239
4.2388 | -84.2401
4.2392 | -84.2404
4.2395 | -84.241
4.2398 | -84.2417
4.2401 | -84.2416
4.2404 | -84.2418
4.2407 | -84.2416
4.241 | -84.2416
4.2413 | -84.2417
4.2416 | -84.2427
4.2419 | -84.2442
4.2423 | -84.2447
4.2426 | -84.2455
4.2429 | -84.2468
4.2432 | -84.2476
4.2435 | -84.2489
4.2438 | -84.2485
4.2441 | -84.2489
4.2444 | -84.2496
4.2447 | -84.2507
4.2451 | -84.2507
4.2454 | -84.2517
4.2457 | -84.2525
4.246 | -84.2528
4.2463 | -84.2537
4.2466 | -84.2546
4.2469 | -84.2557
4.2472 | -84.2555
4.2475 | -84.2557
4.2478 | -84.2565
4.2482 | -84.2574
4.2485 | -84.2577
4.2488 | -84.258
4.2491 | -84.2595
4.2494 | -84.2605
4.2497 | -84.2603
4.25 | -84.2607
4.2503 | -84.2608
4.2506 | -84.261
4.251 | -84.2614
4.2513 | -84.2612
4.2516 | -84.2606
4.2519 | -84.2597
4.2522 | -84.2597
4.2525 | -84.2605
4.2528 | -84.2614
4.2531 | -84.2616
4.2534 | -84.2612
4.2537 | -84.2627
4.2541 | -84.264
4.2544 | -84.2646
4.2547 | -84.2672
4.255 | -84.2713
4.2553 | -84.273
4.2556 | -84.2744
4.2559 | -84.275
4.2562 | -84.276
4.2565 | -84.2759
4.2569 | -84.2758
4.2572 | -84.2774
4.2575 | -84.2792
4.2578 | -84.2796
4.2581 | -84.2802
4.2584 | -84.2806
4.2587 | -84.2805
4.259 | -84.2805
4.2593 | -84.281
4.2596 | -84.2814
4.26 | -84.2822
4.2603 | -84.2831
4.2606 | -84.2837
4.2609 | -84.2841
4.2612 | -84.2848
4.2615 | -84.2856
4.2618 | -84.2861
4.2621 | -84.2857
4.2624 | -84.2872
4.2628 | -84.2895
4.2631 | -84.2914
4.2634 | -84.2926
4.2637 | -84.2934
4.264 | -84.2939
4.2643 | -84.2933
4.2646 | -84.294
4.2649 | -84.295
4.2652 | -84.2956
4.2655 | -84.2964
4.2659 | -84.2972
4.2662 | -84.2983
4.2665 | -84.2995
4.2668 | -84.3002
4.2671 | -84.3011
4.2674 | -84.3009
4.2677 | -84.3009
4.268 | -84.3004
4.2683 | -84.2993
4.2687 | -84.2996
4.269 | -84.3014
4.2693 | -84.3024
4.2696 | -84.3032
4.2699 | -84.3033
4.2702 | -84.303
4.2705 | -84.3016
4.2708 | -84.3017
4.2711 | -84.303
4.2714 | -84.3044
4.2718 | -84.3051
4.2721 | -84.306
4.2724 | -84.3075
4.2727 | -84.3093
4.273 | -84.3096
4.2733 | -84.3091
4.2736 | -84.3084
4.2739 | -84.3093
4.2742 | -84.3092
4.2746 | -84.3102
4.2749 | -84.3118
4.2752 | -84.3131
4.2755 | -84.3135
4.2758 | -84.3145
4.2761 | -84.3163
4.2764 | -84.3167
4.2767 | -84.3168
4.277 | -84.3166
4.2773 | -84.3172
4.2777 | -84.3184
4.278 | -84.3198
4.2783 | -84.3214
4.2786 | -84.3225
4.2789 | -84.3237
4.2792 | -84.3241
4.2795 | -84.3254
4.2798 | -84.3276
4.2801 | -84.3293
4.2805 | -84.3303
4.2808 | -84.331
4.2811 | -84.3314
4.2814 | -84.3321
4.2817 | -84.3329
4.282 | -84.3338
4.2823 | -84.334
4.2826 | -84.3351
4.2829 | -84.3361
4.2832 | -84.3373
4.2836 | -84.3382
4.2839 | -84.3374
4.2842 | -84.3372
4.2845 | -84.3376
4.2848 | -84.3381
4.2851 | -84.34
4.2854 | -84.341
4.2857 | -84.3416
4.286 | -84.3431
4.2864 | -84.3442
4.2867 | -84.3461
4.287 | -84.3476
4.2873 | -84.3488
4.2876 | -84.3493
4.2879 | -84.3505
4.2882 | -84.3517
4.2885 | -84.3529
4.2888 | -84.355
4.2891 | -84.3565
4.2895 | -84.3573
4.2898 | -84.3588
4.2901 | -84.3608
4.2904 | -84.3612
4.2907 | -84.3621
4.291 | -84.3629
4.2913 | -84.363
4.2916 | -84.3645
4.2919 | -84.3656
4.2923 | -84.3666
4.2926 | -84.3673
4.2929 | -84.3686
4.2932 | -84.3699
4.2935 | -84.3715
4.2938 | -84.3727
4.2941 | -84.3743
4.2944 | -84.3776
4.2947 | -84.3804
4.295 | -84.3826
4.2954 | -84.3847
4.2957 | -84.386
4.296 | -84.3883
4.2963 | -84.3907
4.2966 | -84.3925
4.2969 | -84.394
4.2972 | -84.3953
4.2975 | -84.3965
4.2978 | -84.3966
4.2982 | -84.3976
4.2985 | -84.3986
4.2988 | -84.3993
4.2991 | -84.4004
4.2994 | -84.4011
4.2997 | -84.4025
4.3 | -84.4049
4.3003 | -84.4063
4.3006 | -84.4072
4.3009 | -84.4055
4.3013 | -84.4049
4.3016 | -84.4066
4.3019 | -84.4077
4.3022 | -84.4079
4.3025 | -84.4092
4.3028 | -84.4104
4.3031 | -84.412
4.3034 | -84.4145
4.3037 | -84.4167
4.3041 | -84.4179
4.3044 | -84.4191
4.3047 | -84.4202
4.305 | -84.4212
4.3053 | -84.4232
4.3056 | -84.4245
4.3059 | -84.4257
4.3062 | -84.4272
4.3065 | -84.4281
4.3068 | -84.4296
4.3072 | -84.4303
4.3075 | -84.4304
4.3078 | -84.4322
4.3081 | -84.4335
4.3084 | -84.434
4.3087 | -84.4347
4.309 | -84.4355
4.3093 | -84.4367
4.3096 | -84.4378
4.31 | -84.4379
4.3103 | -84.4391
4.3106 | -84.44
4.3109 | -84.4411
4.3112 | -84.4419
4.3115 | -84.443
4.3118 | -84.4445
4.3121 | -84.446
4.3124 | -84.4482
4.3127 | -84.4503
4.3131 | -84.4516
4.3134 | -84.4525
4.3137 | -84.4528
4.314 | -84.4539
4.3143 | -84.4549
4.3146 | -84.4553
4.3149 | -84.4565
4.3152 | -84.4575
4.3155 | -84.4582
4.3159 | -84.4596
4.3162 | -84.4606
4.3165 | -84.4615
4.3168 | -84.4621
4.3171 | -84.464
4.3174 | -84.4662
4.3177 | -84.467
4.318 | -84.4674
4.3183 | -84.4689
4.3186 | -84.4703
4.319 | -84.4724
4.3193 | -84.4763
4.3196 | -84.4766
4.3199 | -84.4779
4.3202 | -84.4807
4.3205 | -84.4823
4.3208 | -84.4836
4.3211 | -84.4852
4.3214 | -84.4861
4.3218 | -84.4867
4.3221 | -84.4877
4.3224 | -84.4895
4.3227 | -84.4909
4.323 | -84.4928
4.3233 | -84.4938
4.3236 | -84.495
4.3239 | -84.4965
4.3242 | -84.4979
4.3245 | -84.4992
4.3249 | -84.5004
4.3252 | -84.5017
4.3255 | -84.5034
4.3258 | -84.5044
4.3261 | -84.5047
4.3264 | -84.5055
4.3267 | -84.5073
4.327 | -84.5092
4.3273 | -84.5106
4.3277 | -84.5121
4.328 | -84.5136
4.3283 | -84.5155
4.3286 | -84.5166
4.3289 | -84.5179
4.3292 | -84.5192
4.3295 | -84.5199
4.3298 | -84.5212
4.3301 | -84.5226
4.3304 | -84.5236
4.3308 | -84.5251
4.3311 | -84.527
4.3314 | -84.5281
4.3317 | -84.5298
4.332 | -84.5324
4.3323 | -84.5352
4.3326 | -84.5348
4.3329 | -84.5349
4.3332 | -84.538
4.3336 | -84.5408
4.3339 | -84.5416
4.3342 | -84.5432
4.3345 | -84.5442
4.3348 | -84.5454
4.3351 | -84.5467
4.3354 | -84.5479
4.3357 | -84.5496
4.336 | -84.5509
4.3364 | -84.5527
4.3367 | -84.5535
4.337 | -84.5543
4.3373 | -84.5556
4.3376 | -84.5566
4.3379 | -84.5572
4.3382 | -84.5595
4.3385 | -84.5619
4.3388 | -84.5623
4.3391 | -84.5619
4.3395 | -84.5626
4.3398 | -84.5641
4.3401 | -84.5669
4.3404 | -84.5698
4.3407 | -84.5718
4.341 | -84.5726
4.3413 | -84.5734
4.3416 | -84.5758
4.3419 | -84.5771
4.3422 | -84.5775
4.3426 | -84.5799
4.3429 | -84.5817
4.3432 | -84.582
4.3435 | -84.5844
4.3438 | -84.5873
4.3441 | -84.5898
4.3444 | -84.5917
4.3447 | -84.5929
4.345 | -84.5941
4.3454 | -84.5945
4.3457 | -84.5946
4.346 | -84.5958
4.3463 | -84.5971
4.3466 | -84.5985
4.3469 | -84.5992
4.3472 | -84.6001
4.3475 | -84.6015
4.3478 | -84.6034
4.3482 | -84.6038
4.3485 | -84.6048
4.3488 | -84.6056
4.3491 | -84.6065
4.3494 | -84.609
4.3497 | -84.6115
4.35 | -84.6129
4.3503 | -84.6142
4.3506 | -84.615
4.3509 | -84.6157
4.3513 | -84.6179
4.3516 | -84.6194
4.3519 | -84.62
4.3522 | -84.6221
4.3525 | -84.6238
4.3528 | -84.6249
4.3531 | -84.6259
4.3534 | -84.6277
4.3537 | -84.6307
4.354 | -84.6327
4.3544 | -84.6332
4.3547 | -84.6336
4.355 | -84.6347
4.3553 | -84.6358
4.3556 | -84.6366
4.3559 | -84.6372
4.3562 | -84.6391
4.3565 | -84.6394
4.3568 | -84.6403
4.3572 | -84.6415
4.3575 | -84.6432
4.3578 | -84.6451
4.3581 | -84.6469
4.3584 | -84.6487
4.3587 | -84.6503
4.359 | -84.6516
4.3593 | -84.6536
4.3596 | -84.6547
4.36 | -84.6564
4.3603 | -84.6577
4.3606 | -84.6586
4.3609 | -84.6598
4.3612 | -84.6604
4.3615 | -84.6604
4.3618 | -84.6615
4.3621 | -84.6625
4.3624 | -84.6638
4.3627 | -84.6653
4.3631 | -84.6665
4.3634 | -84.668
4.3637 | -84.6684
4.364 | -84.6685
4.3643 | -84.6695
4.3646 | -84.6707
4.3649 | -84.6716
4.3652 | -84.6731
4.3655 | -84.6741
4.3659 | -84.6749
4.3662 | -84.677
4.3665 | -84.6803
4.3668 | -84.6824
4.3671 | -84.6835
4.3674 | -84.6861
4.3677 | -84.6882
4.368 | -84.6894
4.3683 | -84.6903
4.3686 | -84.6909
4.369 | -84.6918
4.3693 | -84.6928
4.3696 | -84.6957
4.3699 | -84.6986
4.3702 | -84.6988
4.3705 | -84.6996
4.3708 | -84.7007
4.3711 | -84.7021
4.3714 | -84.7034
4.3718 | -84.7047
4.3721 | -84.7043
4.3724 | -84.7053
4.3727 | -84.7062
4.373 | -84.7072
4.3733 | -84.7082
4.3736 | -84.7104
4.3739 | -84.7117
4.3742 | -84.7124
4.3745 | -84.7137
4.3749 | -84.7145
4.3752 | -84.7149
4.3755 | -84.7159
4.3758 | -84.7169
4.3761 | -84.7181
4.3764 | -84.7192
4.3767 | -84.7201
4.377 | -84.7204
4.3773 | -84.721
4.3777 | -84.7217
4.378 | -84.7228
4.3783 | -84.7237
4.3786 | -84.7254
4.3789 | -84.7264
4.3792 | -84.7282
4.3795 | -84.7308
4.3798 | -84.7323
4.3801 | -84.733
4.3804 | -84.7337
4.3808 | -84.7345
4.3811 | -84.736
4.3814 | -84.7381
4.3817 | -84.7406
4.382 | -84.7424
4.3823 | -84.7433
4.3826 | -84.7451
4.3829 | -84.7462
4.3832 | -84.7473
4.3836 | -84.7485
4.3839 | -84.7482
4.3842 | -84.7492
4.3845 | -84.7496
4.3848 | -84.7504
4.3851 | -84.7515
4.3854 | -84.7521
4.3857 | -84.7524
4.386 | -84.7532
4.3863 | -84.7535
4.3867 | -84.7544
4.387 | -84.7558
4.3873 | -84.7569
4.3876 | -84.7579
4.3879 | -84.7587
4.3882 | -84.7606
4.3885 | -84.7644
4.3888 | -84.7655
4.3891 | -84.7661
4.3895 | -84.7659
4.3898 | -84.7666
4.3901 | -84.768
4.3904 | -84.7684
4.3907 | -84.7689
4.391 | -84.7692
4.3913 | -84.7704
4.3916 | -84.7726
4.3919 | -84.7745
4.3922 | -84.7751
4.3926 | -84.7758
4.3929 | -84.7763
4.3932 | -84.7756
4.3935 | -84.7758
4.3938 | -84.7767
4.3941 | -84.7776
4.3944 | -84.7788
4.3947 | -84.7803
4.395 | -84.781
4.3954 | -84.7834
4.3957 | -84.7871
4.396 | -84.7888
4.3963 | -84.7906
4.3966 | -84.7914
4.3969 | -84.792
4.3972 | -84.7927
4.3975 | -84.7932
4.3978 | -84.7938
4.3981 | -84.7948
4.3985 | -84.7961
4.3988 | -84.7967
4.3991 | -84.7985
4.3994 | -84.7992
4.3997 | -84.8
4.4 | -84.8012
4.4003 | -84.8031
4.4006 | -84.8062
4.4009 | -84.8078
4.4013 | -84.8088
4.4016 | -84.8097
4.4019 | -84.8105
4.4022 | -84.8114
4.4025 | -84.8124
4.4028 | -84.8143
4.4031 | -84.8123
4.4034 | -84.8121
4.4037 | -84.8137
4.404 | -84.8146
4.4044 | -84.8158
4.4047 | -84.816
4.405 | -84.8167
4.4053 | -84.8177
4.4056 | -84.8182
4.4059 | -84.8168
4.4062 | -84.8163
4.4065 | -84.8171
4.4068 | -84.8181
4.4072 | -84.8189
4.4075 | -84.8195
4.4078 | -84.82
4.4081 | -84.8216
4.4084 | -84.8227
4.4087 | -84.8241
4.409 | -84.8258
4.4093 | -84.827
4.4096 | -84.8278
4.4099 | -84.829
4.4103 | -84.8289
4.4106 | -84.8284
4.4109 | -84.8288
4.4112 | -84.8304
4.4115 | -84.8312
4.4118 | -84.8325
4.4121 | -84.8336
4.4124 | -84.8343
4.4127 | -84.8348
4.4131 | -84.8378
4.4134 | -84.8387
4.4137 | -84.8381
4.414 | -84.8384
4.4143 | -84.8394
4.4146 | -84.8407
4.4149 | -84.8411
4.4152 | -84.8416
4.4155 | -84.8432
4.4158 | -84.8437
4.4162 | -84.8448
4.4165 | -84.8464
4.4168 | -84.8484
4.4171 | -84.8493
4.4174 | -84.8486
4.4177 | -84.8496
4.418 | -84.8493
4.4183 | -84.8497
4.4186 | -84.8499
4.419 | -84.8497
4.4193 | -84.8492
4.4196 | -84.8503
4.4199 | -84.8514
4.4202 | -84.8506
4.4205 | -84.8498
4.4208 | -84.849
4.4211 | -84.8501
4.4214 | -84.851
4.4217 | -84.8522
4.4221 | -84.8537
4.4224 | -84.8539
4.4227 | -84.8549
4.423 | -84.8562
4.4233 | -84.8576
4.4236 | -84.8592
4.4239 | -84.8602
4.4242 | -84.861
4.4245 | -84.8616
4.4249 | -84.8616
4.4252 | -84.8623
4.4255 | -84.863
4.4258 | -84.8638
4.4261 | -84.8653
4.4264 | -84.8664
4.4267 | -84.8669
4.427 | -84.8677
4.4273 | -84.8697
4.4276 | -84.8725
4.428 | -84.8742
4.4283 | -84.8754
4.4286 | -84.8765
4.4289 | -84.8772
4.4292 | -84.8791
4.4295 | -84.882
4.4298 | -84.8817
4.4301 | -84.8823
4.4304 | -84.8832
4.4308 | -84.883
4.4311 | -84.884
4.4314 | -84.8855
4.4317 | -84.8871
4.432 | -84.8866
4.4323 | -84.8852
4.4326 | -84.8852
4.4329 | -84.8862
4.4332 | -84.8874
4.4335 | -84.8886
4.4339 | -84.8894
4.4342 | -84.8899
4.4345 | -84.8909
4.4348 | -84.8922
4.4351 | -84.8935
4.4354 | -84.8955
4.4357 | -84.8972
4.436 | -84.9002
4.4363 | -84.9017
4.4367 | -84.9017
4.437 | -84.9055
4.4373 | -84.9079
4.4376 | -84.9094
4.4379 | -84.9099
4.4382 | -84.9118
4.4385 | -84.9131
4.4388 | -84.9139
4.4391 | -84.9156
4.4394 | -84.9163
4.4398 | -84.9165
4.4401 | -84.9163
4.4404 | -84.9179
4.4407 | -84.9185
4.441 | -84.9192
4.4413 | -84.9204
4.4416 | -84.9214
4.4419 | -84.9227
4.4422 | -84.9245
4.4426 | -84.9263
4.4429 | -84.9271
4.4432 | -84.9279
4.4435 | -84.9277
4.4438 | -84.9291
4.4441 | -84.9326
4.4444 | -84.9368
4.4447 | -84.9385
4.445 | -84.9392
4.4453 | -84.9397
4.4457 | -84.9407
4.446 | -84.942
4.4463 | -84.9429
4.4466 | -84.9433
4.4469 | -84.9433
4.4472 | -84.9416
4.4475 | -84.9416
4.4478 | -84.9415
4.4481 | -84.9415
4.4485 | -84.9421
4.4488 | -84.9427
4.4491 | -84.943
4.4494 | -84.943
4.4497 | -84.9438
4.45 | -84.9449
4.4503 | -84.9462
4.4506 | -84.9472
4.4509 | -84.9485
4.4512 | -84.9493
4.4516 | -84.9498
4.4519 | -84.9508
4.4522 | -84.9528
4.4525 | -84.9537
4.4528 | -84.9557
4.4531 | -84.9567
4.4534 | -84.9571
4.4537 | -84.9572
4.454 | -84.9587
4.4544 | -84.9594
4.4547 | -84.9602
4.455 | -84.96
4.4553 | -84.9608
4.4556 | -84.9618
4.4559 | -84.9619
4.4562 | -84.9625
4.4565 | -84.9634
4.4568 | -84.9642
4.4571 | -84.9651
4.4575 | -84.966
4.4578 | -84.967
4.4581 | -84.9675
4.4584 | -84.9693
4.4587 | -84.9716
4.459 | -84.9722
4.4593 | -84.9737
4.4596 | -84.9748
4.4599 | -84.976
4.4603 | -84.9782
4.4606 | -84.9796
4.4609 | -84.9806
4.4612 | -84.9814
4.4615 | -84.9821
4.4618 | -84.9826
4.4621 | -84.9837
4.4624 | -84.9851
4.4627 | -84.986
4.463 | -84.9869
4.4634 | -84.9871
4.4637 | -84.9882
4.464 | -84.9904
4.4643 | -84.9906
4.4646 | -84.9914
4.4649 | -84.9927
4.4652 | -84.9965
4.4655 | -85.0019
4.4658 | -85.005
4.4662 | -85.0061
4.4665 | -85.0068
4.4668 | -85.0073
4.4671 | -85.0078
4.4674 | -85.0075
4.4677 | -85.0088
4.468 | -85.0099
4.4683 | -85.0111
4.4686 | -85.0114
4.4689 | -85.0118
4.4693 | -85.0133
4.4696 | -85.0149
4.4699 | -85.0179
4.4702 | -85.0193
4.4705 | -85.0189
4.4708 | -85.0198
4.4711 | -85.0208
4.4714 | -85.0212
4.4717 | -85.0214
4.4721 | -85.0218
4.4724 | -85.0216
4.4727 | -85.022
4.473 | -85.0226
4.4733 | -85.0246
4.4736 | -85.026
4.4739 | -85.0269
4.4742 | -85.0314
4.4745 | -85.0352
4.4748 | -85.0359
4.4752 | -85.0364
4.4755 | -85.0364
4.4758 | -85.0375
4.4761 | -85.0386
4.4764 | -85.04
4.4767 | -85.0409
4.477 | -85.043
4.4773 | -85.0438
4.4776 | -85.0444
4.478 | -85.0444
4.4783 | -85.0445
4.4786 | -85.0456
4.4789 | -85.0461
4.4792 | -85.0475
4.4795 | -85.0486
4.4798 | -85.0496
4.4801 | -85.0512
4.4804 | -85.0532
4.4807 | -85.0547
4.4811 | -85.0557
4.4814 | -85.0574
4.4817 | -85.0579
4.482 | -85.0585
4.4823 | -85.0597
4.4826 | -85.0604
4.4829 | -85.0603
4.4832 | -85.0603
4.4835 | -85.0608
4.4839 | -85.0607
4.4842 | -85.0608
4.4845 | -85.0624
4.4848 | -85.0637
4.4851 | -85.0643
4.4854 | -85.0657
4.4857 | -85.0671
4.486 | -85.067
4.4863 | -85.0673
4.4867 | -85.0689
4.487 | -85.0688
4.4873 | -85.0707
4.4876 | -85.0731
4.4879 | -85.074
4.4882 | -85.0747
4.4885 | -85.0748
4.4888 | -85.0707
4.4891 | -85.071
4.4894 | -85.0715
4.4898 | -85.072
4.4901 | -85.0726
4.4904 | -85.0723
4.4907 | -85.0727
4.491 | -85.0748
4.4913 | -85.0752
4.4916 | -85.0761
4.4919 | -85.0767
4.4922 | -85.0769
4.4925 | -85.0777
4.4929 | -85.0798
4.4932 | -85.0807
4.4935 | -85.0811
4.4938 | -85.0814
4.4941 | -85.0814
4.4944 | -85.082
4.4947 | -85.0822
4.495 | -85.0841
4.4953 | -85.0862
4.4957 | -85.087
4.496 | -85.0877
4.4963 | -85.0881
4.4966 | -85.0886
4.4969 | -85.089
4.4972 | -85.0898
4.4975 | -85.0901
4.4978 | -85.0898
4.4981 | -85.0894
4.4985 | -85.0896
4.4988 | -85.0908
4.4991 | -85.0917
4.4994 | -85.0926
4.4997 | -85.0931
4.5 | -85.0928
4.5003 | -85.0931
4.5006 | -85.0964
4.5009 | -85.0977
4.5012 | -85.0982
4.5016 | -85.0982
4.5019 | -85.0982
4.5022 | -85.0989
4.5025 | -85.0997
4.5028 | -85.0997
4.5031 | -85.0993
4.5034 | -85.0996
4.5037 | -85.1001
4.504 | -85.1015
4.5044 | -85.1022
4.5047 | -85.1026
4.505 | -85.1043
4.5053 | -85.1049
4.5056 | -85.1053
4.5059 | -85.1058
4.5062 | -85.106
4.5065 | -85.1056
4.5068 | -85.1055
4.5071 | -85.1059
4.5075 | -85.1057
4.5078 | -85.1057
4.5081 | -85.1052
4.5084 | -85.1051
4.5087 | -85.1054
4.509 | -85.106
4.5093 | -85.1053
4.5096 | -85.1049
4.5099 | -85.1057
4.5103 | -85.1066
4.5106 | -85.1067
4.5109 | -85.1068
4.5112 | -85.1072
4.5115 | -85.1077
4.5118 | -85.1082
4.5121 | -85.1089
4.5124 | -85.1099
4.5127 | -85.1104
4.513 | -85.1101
4.5134 | -85.1104
4.5137 | -85.111
4.514 | -85.1114
4.5143 | -85.1126
4.5146 | -85.1136
4.5149 | -85.114
4.5152 | -85.1147
4.5155 | -85.1146
4.5158 | -85.1156
4.5162 | -85.1149
4.5165 | -85.113
4.5168 | -85.1125
4.5171 | -85.1136
4.5174 | -85.1142
4.5177 | -85.1141
4.518 | -85.1132
4.5183 | -85.1124
4.5186 | -85.1121
4.5189 | -85.1128
4.5193 | -85.1141
4.5196 | -85.1153
4.5199 | -85.1157
4.5202 | -85.1154
4.5205 | -85.1157
4.5208 | -85.1156
4.5211 | -85.1156
4.5214 | -85.1161
4.5217 | -85.1164
4.5221 | -85.1167
4.5224 | -85.1163
4.5227 | -85.1146
4.523 | -85.1134
4.5233 | -85.1124
4.5236 | -85.1103
4.5239 | -85.1091
4.5242 | -85.1091
4.5245 | -85.104
4.5248 | -85.1023
4.5252 | -85.101
4.5255 | -85.1007
4.5258 | -85.1008
4.5261 | -85.1008
4.5264 | -85.1014
4.5267 | -85.1019
4.527 | -85.1019
4.5273 | -85.1027
4.5276 | -85.1021
4.528 | -85.1006
4.5283 | -85.1003
4.5286 | -85.1007
4.5289 | -85.1011
4.5292 | -85.1009
4.5295 | -85.1008
4.5298 | -85.1003
4.5301 | -85.0998
4.5304 | -85.1002
4.5307 | -85.1005
4.5311 | -85.1007
4.5314 | -85.1007
4.5317 | -85.1005
4.532 | -85.1003
4.5323 | -85.0999
4.5326 | -85.0994
4.5329 | -85.099
4.5332 | -85.0979
4.5335 | -85.0958
4.5339 | -85.0956
4.5342 | -85.096
4.5345 | -85.0968
4.5348 | -85.0967
4.5351 | -85.0963
4.5354 | -85.0953
4.5357 | -85.0952
4.536 | -85.0951
4.5363 | -85.095
4.5366 | -85.0949
4.537 | -85.0949
4.5373 | -85.0947
4.5376 | -85.0944
4.5379 | -85.0935
4.5382 | -85.0928
4.5385 | -85.0929
4.5388 | -85.0936
4.5391 | -85.0927
4.5394 | -85.0939
4.5398 | -85.0951
4.5401 | -85.0949
4.5404 | -85.0949
4.5407 | -85.095
4.541 | -85.0951
4.5413 | -85.0948
4.5416 | -85.0942
4.5419 | -85.094
4.5422 | -85.0943
4.5425 | -85.094
4.5429 | -85.0955
4.5432 | -85.0968
4.5435 | -85.0964
4.5438 | -85.0953
4.5441 | -85.0941
4.5444 | -85.0938
4.5447 | -85.0933
4.545 | -85.0924
4.5453 | -85.0926
4.5457 | -85.0915
4.546 | -85.0902
4.5463 | -85.0903
4.5466 | -85.0906
4.5469 | -85.0906
4.5472 | -85.0891
4.5475 | -85.0866
4.5478 | -85.0846
4.5481 | -85.0836
4.5484 | -85.0824
4.5488 | -85.0813
4.5491 | -85.0821
4.5494 | -85.0817
4.5497 | -85.0812
4.55 | -85.0819
4.5503 | -85.0825
4.5506 | -85.0822
4.5509 | -85.0814
4.5512 | -85.0809
4.5516 | -85.08
4.5519 | -85.0796
4.5522 | -85.0793
4.5525 | -85.0793
4.5528 | -85.0791
4.5531 | -85.0794
4.5534 | -85.0793
4.5537 | -85.0789
4.554 | -85.0783
4.5543 | -85.0783
4.5547 | -85.0777
4.555 | -85.0767
4.5553 | -85.0763
4.5556 | -85.0764
4.5559 | -85.0753
4.5562 | -85.0742
4.5565 | -85.074
4.5568 | -85.0732
4.5571 | -85.0737
4.5575 | -85.0745
4.5578 | -85.0755
4.5581 | -85.0746
4.5584 | -85.0733
4.5587 | -85.0728
4.559 | -85.0724
4.5593 | -85.0731
4.5596 | -85.0715
4.5599 | -85.0703
4.5602 | -85.0697
4.5606 | -85.0682
4.5609 | -85.0685
4.5612 | -85.0684
4.5615 | -85.0668
4.5618 | -85.0662
4.5621 | -85.0658
4.5624 | -85.0654
4.5627 | -85.0647
4.563 | -85.0642
4.5634 | -85.0638
4.5637 | -85.0631
4.564 | -85.0626
4.5643 | -85.0622
4.5646 | -85.0614
4.5649 | -85.0604
4.5652 | -85.0606
4.5655 | -85.06
4.5658 | -85.06
4.5661 | -85.0595
4.5665 | -85.0593
4.5668 | -85.0585
4.5671 | -85.058
4.5674 | -85.0575
4.5677 | -85.057
4.568 | -85.0563
4.5683 | -85.0549
4.5686 | -85.0543
4.5689 | -85.0534
4.5693 | -85.0518
4.5696 | -85.0513
4.5699 | -85.0507
4.5702 | -85.0506
4.5705 | -85.05
4.5708 | -85.0498
4.5711 | -85.0508
4.5714 | -85.0517
4.5717 | -85.0506
4.572 | -85.0492
4.5724 | -85.0477
4.5727 | -85.0468
4.573 | -85.046
4.5733 | -85.0452
4.5736 | -85.0452
4.5739 | -85.0449
4.5742 | -85.0445
4.5745 | -85.0442
4.5748 | -85.0436
4.5752 | -85.044
4.5755 | -85.0452
4.5758 | -85.0453
4.5761 | -85.0448
4.5764 | -85.0444
4.5767 | -85.0438
4.577 | -85.0437
4.5773 | -85.0437
4.5776 | -85.0442
4.5779 | -85.0441
4.5783 | -85.043
4.5786 | -85.0422
4.5789 | -85.0419
4.5792 | -85.0395
4.5795 | -85.0382
4.5798 | -85.0372
4.5801 | -85.0356
4.5804 | -85.0344
4.5807 | -85.0345
4.5811 | -85.0344
4.5814 | -85.0333
4.5817 | -85.0315
4.582 | -85.0308
4.5823 | -85.0289
4.5826 | -85.0256
4.5829 | -85.0248
4.5832 | -85.024
4.5835 | -85.0236
4.5838 | -85.0228
4.5842 | -85.0226
4.5845 | -85.0216
4.5848 | -85.022
4.5851 | -85.0219
4.5854 | -85.0212
4.5857 | -85.0205
4.586 | -85.019
4.5863 | -85.0171
4.5866 | -85.0159
4.587 | -85.0148
4.5873 | -85.0146
4.5876 | -85.0114
4.5879 | -85.0085
4.5882 | -85.0073
4.5885 | -85.0064
4.5888 | -85.0058
4.5891 | -85.0053
4.5894 | -85.0048
4.5897 | -85.0042
4.5901 | -85.0034
4.5904 | -85.0035
4.5907 | -85.0033
4.591 | -85.0029
4.5913 | -85.0024
4.5916 | -85.0014
4.5919 | -85.0001
4.5922 | -85.0009
4.5925 | -85.0014
4.5929 | -85.0013
4.5932 | -85.0013
4.5935 | -85.0006
4.5938 | -85
4.5941 | -84.9996
4.5944 | -84.999
4.5947 | -84.9996
4.595 | -84.9997
4.5953 | -84.999
4.5956 | -84.9984
4.596 | -84.9978
4.5963 | -84.9963
4.5966 | -84.9957
4.5969 | -84.9948
4.5972 | -84.994
4.5975 | -84.9932
4.5978 | -84.9924
4.5981 | -84.9919
4.5984 | -84.9912
4.5988 | -84.9908
4.5991 | -84.9902
4.5994 | -84.988
4.5997 | -84.987
4.6 | -84.9876
4.6003 | -84.9873
4.6006 | -84.9862
4.6009 | -84.9851
4.6012 | -84.9845
4.6015 | -84.9844
4.6019 | -84.9838
4.6022 | -84.9834
4.6025 | -84.9823
4.6028 | -84.9815
4.6031 | -84.9805
4.6034 | -84.9791
4.6037 | -84.9781
4.604 | -84.9774
4.6043 | -84.9767
4.6047 | -84.9759
4.605 | -84.975
4.6053 | -84.974
4.6056 | -84.9726
4.6059 | -84.9719
4.6062 | -84.9712
4.6065 | -84.9704
4.6068 | -84.9694
4.6071 | -84.9696
4.6074 | -84.9687
4.6078 | -84.9669
4.6081 | -84.9663
4.6084 | -84.9655
4.6087 | -84.9647
4.609 | -84.9637
4.6093 | -84.9627
4.6096 | -84.9621
4.6099 | -84.9616
4.6102 | -84.9618
4.6106 | -84.9645
4.6109 | -84.9655
4.6112 | -84.9647
4.6115 | -84.9643
4.6118 | -84.9646
4.6121 | -84.9639
4.6124 | -84.963
4.6127 | -84.9628
4.613 | -84.9622
4.6133 | -84.9608
4.6137 | -84.9595
4.614 | -84.9584
4.6143 | -84.958
4.6146 | -84.9573
4.6149 | -84.9566
4.6152 | -84.9556
4.6155 | -84.9548
4.6158 | -84.9534
4.6161 | -84.9527
4.6165 | -84.9524
4.6168 | -84.9509
4.6171 | -84.9508
4.6174 | -84.9516
4.6177 | -84.9511
4.618 | -84.9503
4.6183 | -84.9499
4.6186 | -84.9489
4.6189 | -84.9483
4.6193 | -84.9477
4.6196 | -84.9472
4.6199 | -84.9476
4.6202 | -84.9484
4.6205 | -84.9478
4.6208 | -84.9474
4.6211 | -84.9454
4.6214 | -84.9447
4.6217 | -84.9438
4.622 | -84.9426
4.6224 | -84.9423
4.6227 | -84.9424
4.623 | -84.9426
4.6233 | -84.9421
4.6236 | -84.9412
4.6239 | -84.942
4.6242 | -84.9423
4.6245 | -84.9413
4.6248 | -84.9416
4.6251 | -84.9421
4.6255 | -84.9414
4.6258 | -84.9411
4.6261 | -84.941
4.6264 | -84.941
4.6267 | -84.9413
4.627 | -84.9416
4.6273 | -84.9417
4.6276 | -84.9411
4.6279 | -84.9401
4.6283 | -84.9401
4.6286 | -84.9401
4.6289 | -84.9398
4.6292 | -84.9392
4.6295 | -84.9386
4.6298 | -84.9383
4.6301 | -84.938
4.6304 | -84.9389
4.6307 | -84.9384
4.6311 | -84.9377
4.6314 | -84.9367
4.6317 | -84.9355
4.632 | -84.9359
4.6323 | -84.9346
4.6326 | -84.9342
4.6329 | -84.9344
4.6332 | -84.9342
4.6335 | -84.9337
4.6338 | -84.9333
4.6342 | -84.933
4.6345 | -84.9329
4.6348 | -84.933
4.6351 | -84.9325
4.6354 | -84.9322
4.6357 | -84.9311
4.636 | -84.9297
4.6363 | -84.9293
4.6366 | -84.9288
4.6369 | -84.9287
4.6373 | -84.9293
4.6376 | -84.9297
4.6379 | -84.9291
4.6382 | -84.9285
4.6385 | -84.9266
4.6388 | -84.9258
4.6391 | -84.9252
4.6394 | -84.9246
4.6397 | -84.9244
4.6401 | -84.9245
4.6404 | -84.924
4.6407 | -84.924
4.641 | -84.9241
4.6413 | -84.924
4.6416 | -84.9241
4.6419 | -84.9235
4.6422 | -84.9221
4.6425 | -84.9219
4.6429 | -84.9215
4.6432 | -84.9212
4.6435 | -84.9209
4.6438 | -84.9205
4.6441 | -84.9203
4.6444 | -84.9222
4.6447 | -84.9235
4.645 | -84.9231
4.6453 | -84.9231
4.6456 | -84.923
4.646 | -84.9223
4.6463 | -84.9224
4.6466 | -84.9224
4.6469 | -84.9222
4.6472 | -84.9221
4.6475 | -84.9218
4.6478 | -84.9213
4.6481 | -84.9208
4.6484 | -84.9206
4.6488 | -84.9207
4.6491 | -84.9208
4.6494 | -84.9189
4.6497 | -84.9176
4.65 | -84.9188
4.6503 | -84.9193
4.6506 | -84.9191
4.6509 | -84.9183
4.6512 | -84.9186
4.6515 | -84.9202
4.6519 | -84.9205
4.6522 | -84.9198
4.6525 | -84.9199
4.6528 | -84.9195
4.6531 | -84.9199
4.6534 | -84.9202
4.6537 | -84.9205
4.654 | -84.9197
4.6543 | -84.9197
4.6547 | -84.9203
4.655 | -84.9203
4.6553 | -84.9204
4.6556 | -84.9203
4.6559 | -84.9201
4.6562 | -84.9202
4.6565 | -84.9195
4.6568 | -84.9193
4.6571 | -84.9195
4.6574 | -84.9192
4.6578 | -84.9189
4.6581 | -84.9197
4.6584 | -84.9204
4.6587 | -84.9198
4.659 | -84.9208
4.6593 | -84.9206
4.6596 | -84.92
4.6599 | -84.9202
4.6602 | -84.9205
4.6606 | -84.92
4.6609 | -84.9204
4.6612 | -84.9213
4.6615 | -84.922
4.6618 | -84.9224
4.6621 | -84.9229
4.6624 | -84.9235
4.6627 | -84.9236
4.663 | -84.9236
4.6633 | -84.9237
4.6637 | -84.9247
4.664 | -84.9254
4.6643 | -84.9252
4.6646 | -84.9253
4.6649 | -84.9257
4.6652 | -84.9259
4.6655 | -84.9266
4.6658 | -84.9262
4.6661 | -84.923
4.6665 | -84.9172
4.6668 | -84.9163
4.6671 | -84.9167
4.6674 | -84.9167
4.6677 | -84.9172
4.668 | -84.9177
4.6683 | -84.9181
4.6686 | -84.9185
4.6689 | -84.9186
4.6692 | -84.9193
4.6696 | -84.9201
4.6699 | -84.9205
4.6702 | -84.9216
4.6705 | -84.9216
4.6708 | -84.9211
4.6711 | -84.9213
4.6714 | -84.9218
4.6717 | -84.9215
4.672 | -84.9205
4.6724 | -84.9197
4.6727 | -84.9194
4.673 | -84.9196
4.6733 | -84.9187
4.6736 | -84.9176
4.6739 | -84.9175
4.6742 | -84.9154
4.6745 | -84.9132
4.6748 | -84.9128
4.6751 | -84.9133
4.6755 | -84.913
4.6758 | -84.9124
4.6761 | -84.9117
4.6764 | -84.9112
4.6767 | -84.9122
4.677 | -84.9113
4.6773 | -84.9111
4.6776 | -84.9115
4.6779 | -84.9122
4.6783 | -84.9135
4.6786 | -84.9133
4.6789 | -84.913
4.6792 | -84.9129
4.6795 | -84.9127
4.6798 | -84.9128
4.6801 | -84.9131
4.6804 | -84.9123
4.6807 | -84.9116
4.681 | -84.9117
4.6814 | -84.9117
4.6817 | -84.9112
4.682 | -84.9107
4.6823 | -84.9108
4.6826 | -84.9113
4.6829 | -84.9119
4.6832 | -84.9124
4.6835 | -84.9129
4.6838 | -84.9131
4.6842 | -84.914
4.6845 | -84.9139
4.6848 | -84.913
4.6851 | -84.9136
4.6854 | -84.914
4.6857 | -84.9148
4.686 | -84.9155
4.6863 | -84.916
4.6866 | -84.9167
4.6869 | -84.9175
4.6873 | -84.9178
4.6876 | -84.918
4.6879 | -84.9179
4.6882 | -84.9184
4.6885 | -84.9188
4.6888 | -84.9189
4.6891 | -84.9191
4.6894 | -84.9195
4.6897 | -84.9202
4.6901 | -84.9212
4.6904 | -84.9215
4.6907 | -84.922
4.691 | -84.9217
4.6913 | -84.9215
4.6916 | -84.9209
4.6919 | -84.9208
4.6922 | -84.9215
4.6925 | -84.9222
4.6928 | -84.9225
4.6932 | -84.923
4.6935 | -84.9235
4.6938 | -84.9235
4.6941 | -84.924
4.6944 | -84.9247
4.6947 | -84.9255
4.695 | -84.9259
4.6953 | -84.9255
4.6956 | -84.9242
4.696 | -84.9238
4.6963 | -84.9243
4.6966 | -84.9246
4.6969 | -84.925
4.6972 | -84.9255
4.6975 | -84.926
4.6978 | -84.9256
4.6981 | -84.925
4.6984 | -84.9261
4.6987 | -84.9272
4.6991 | -84.9286
4.6994 | -84.9296
4.6997 | -84.9303
4.7 | -84.9312
4.7003 | -84.9317
4.7006 | -84.932
4.7009 | -84.9332
4.7012 | -84.9338
4.7015 | -84.9333
4.7019 | -84.934
4.7022 | -84.9337
4.7025 | -84.9332
4.7028 | -84.9326
4.7031 | -84.9328
4.7034 | -84.9335
4.7037 | -84.9342
4.704 | -84.9346
4.7043 | -84.9351
4.7046 | -84.9358
4.705 | -84.9366
4.7053 | -84.938
4.7056 | -84.9386
4.7059 | -84.939
4.7062 | -84.9382
4.7065 | -84.9387
4.7068 | -84.9396
4.7071 | -84.9404
4.7074 | -84.941
4.7078 | -84.9423
4.7081 | -84.9435
4.7084 | -84.944
4.7087 | -84.9442
4.709 | -84.9446
4.7093 | -84.9466
4.7096 | -84.9474
4.7099 | -84.9479
4.7102 | -84.9482
4.7105 | -84.9489
4.7109 | -84.9495
4.7112 | -84.9498
4.7115 | -84.9501
4.7118 | -84.95
4.7121 | -84.9505
4.7124 | -84.9509
4.7127 | -84.9523
4.713 | -84.9532
4.7133 | -84.9539
4.7137 | -84.9546
4.714 | -84.9545
4.7143 | -84.9546
4.7146 | -84.9554
4.7149 | -84.9558
4.7152 | -84.9566
4.7155 | -84.9572
4.7158 | -84.9576
4.7161 | -84.9584
4.7164 | -84.9591
4.7168 | -84.9597
4.7171 | -84.9604
4.7174 | -84.9604
4.7177 | -84.9597
4.718 | -84.9589
4.7183 | -84.9596
4.7186 | -84.9601
4.7189 | -84.9606
4.7192 | -84.9611
4.7196 | -84.9618
4.7199 | -84.9647
4.7202 | -84.9656
4.7205 | -84.965
4.7208 | -84.9656
4.7211 | -84.9659
4.7214 | -84.9671
4.7217 | -84.9682
4.722 | -84.9694
4.7223 | -84.9698
4.7227 | -84.9708
4.723 | -84.974
4.7233 | -84.9761
4.7236 | -84.9767
4.7239 | -84.9771
4.7242 | -84.9776
4.7245 | -84.978
4.7248 | -84.9788
4.7251 | -84.9791
4.7255 | -84.9799
4.7258 | -84.9804
4.7261 | -84.9817
4.7264 | -84.9827
4.7267 | -84.9812
4.727 | -84.9815
4.7273 | -84.982
4.7276 | -84.983
4.7279 | -84.9847
4.7282 | -84.986
4.7286 | -84.9866
4.7289 | -84.9875
4.7292 | -84.9881
4.7295 | -84.9884
4.7298 | -84.9891
4.7301 | -84.9903
4.7304 | -84.9907
4.7307 | -84.9919
4.731 | -84.9932
4.7314 | -84.9946
4.7317 | -84.9952
4.732 | -84.9957
4.7323 | -84.9963
4.7326 | -84.9967
4.7329 | -84.9974
4.7332 | -84.9982
4.7335 | -84.9987
4.7338 | -84.998
4.7341 | -84.9983
4.7345 | -84.9989
4.7348 | -84.9992
4.7351 | -85.0001
4.7354 | -85.0022
4.7357 | -85.0025
4.736 | -85.0036
4.7363 | -85.0052
4.7366 | -85.0061
4.7369 | -85.0075
4.7373 | -85.0085
4.7376 | -85.0091
4.7379 | -85.0091
4.7382 | -85.0091
4.7385 | -85.0093
4.7388 | -85.0112
4.7391 | -85.0119
4.7394 | -85.0129
4.7397 | -85.013
4.74 | -85.0132
4.7404 | -85.0136
4.7407 | -85.0144
4.741 | -85.0154
4.7413 | -85.0158
4.7416 | -85.0156
4.7419 | -85.0157
4.7422 | -85.0158
4.7425 | -85.0159
4.7428 | -85.0163
4.7432 | -85.0173
4.7435 | -85.0181
4.7438 | -85.0191
4.7441 | -85.0197
4.7444 | -85.0203
4.7447 | -85.0207
4.745 | -85.0211
4.7453 | -85.0213
4.7456 | -85.0208
4.7459 | -85.021
4.7463 | -85.0213
4.7466 | -85.0214
4.7469 | -85.0215
4.7472 | -85.0209
4.7475 | -85.0203
4.7478 | -85.0204
4.7481 | -85.0207
4.7484 | -85.0208
4.7487 | -85.0217
4.7491 | -85.0226
4.7494 | -85.0233
4.7497 | -85.0245
4.75 | -85.0249
4.7503 | -85.0253
4.7506 | -85.0258
4.7509 | -85.0275
4.7512 | -85.0306
4.7515 | -85.0316
4.7518 | -85.0318
4.7522 | -85.0316
4.7525 | -85.0312
4.7528 | -85.0312
4.7531 | -85.032
4.7534 | -85.0319
4.7537 | -85.032
4.754 | -85.032
4.7543 | -85.0337
4.7546 | -85.0337
4.755 | -85.0333
4.7553 | -85.0335
4.7556 | -85.0356
4.7559 | -85.0357
4.7562 | -85.0324
4.7565 | -85.0307
4.7568 | -85.0287
4.7571 | -85.0271
4.7574 | -85.0267
4.7577 | -85.0262
4.7581 | -85.0264
4.7584 | -85.0273
4.7587 | -85.0276
4.759 | -85.0272
4.7593 | -85.0278
4.7596 | -85.0283
4.7599 | -85.0286
4.7602 | -85.0296
4.7605 | -85.0302
4.7609 | -85.0305
4.7612 | -85.0298
4.7615 | -85.0295
4.7618 | -85.0298
4.7621 | -85.0297
4.7624 | -85.0295
4.7627 | -85.0303
4.763 | -85.0315
4.7633 | -85.0323
4.7637 | -85.0314
4.764 | -85.031
4.7643 | -85.0312
4.7646 | -85.0314
4.7649 | -85.0319
4.7652 | -85.0325
4.7655 | -85.0325
4.7658 | -85.033
4.7661 | -85.0335
4.7664 | -85.0337
4.7668 | -85.0346
4.7671 | -85.0361
4.7674 | -85.039
4.7677 | -85.0397
4.768 | -85.0368
4.7683 | -85.036
4.7686 | -85.035
4.7689 | -85.0352
4.7692 | -85.0357
4.7695 | -85.0364
4.7699 | -85.0369
4.7702 | -85.0372
4.7705 | -85.0373
4.7708 | -85.0369
4.7711 | -85.0362
4.7714 | -85.0353
4.7717 | -85.0353
4.772 | -85.0356
4.7723 | -85.0352
4.7727 | -85.0353
4.773 | -85.0351
4.7733 | -85.0346
4.7736 | -85.0342
4.7739 | -85.0349
4.7742 | -85.0359
4.7745 | -85.0352
4.7748 | -85.0358
4.7751 | -85.0358
4.7755 | -85.0362
4.7758 | -85.037
4.7761 | -85.037
4.7764 | -85.036
4.7767 | -85.0356
4.777 | -85.0356
4.7773 | -85.0353
4.7776 | -85.0344
4.7779 | -85.0344
4.7782 | -85.0346
4.7786 | -85.0341
4.7789 | -85.0339
4.7792 | -85.0341
4.7795 | -85.0341
4.7798 | -85.0346
4.7801 | -85.0336
4.7804 | -85.0331
4.7807 | -85.0326
4.781 | -85.0326
4.7813 | -85.0324
4.7817 | -85.0322
4.782 | -85.0321
4.7823 | -85.0322
4.7826 | -85.0327
4.7829 | -85.0331
4.7832 | -85.0321
4.7835 | -85.0313
4.7838 | -85.0314
4.7841 | -85.0314
4.7845 | -85.0308
4.7848 | -85.0316
4.7851 | -85.0318
4.7854 | -85.0321
4.7857 | -85.0332
4.786 | -85.0332
4.7863 | -85.0327
4.7866 | -85.0323
4.7869 | -85.0321
4.7873 | -85.0321
4.7876 | -85.0321
4.7879 | -85.0326
4.7882 | -85.033
4.7885 | -85.033
4.7888 | -85.0339
4.7891 | -85.0338
4.7894 | -85.0342
4.7897 | -85.034
4.79 | -85.034
4.7904 | -85.0341
4.7907 | -85.0337
4.791 | -85.0339
4.7913 | -85.034
4.7916 | -85.0333
4.7919 | -85.0331
4.7922 | -85.0329
4.7925 | -85.0318
4.7928 | -85.0307
4.7932 | -85.0305
4.7935 | -85.0306
4.7938 | -85.0308
4.7941 | -85.0312
4.7944 | -85.0321
4.7947 | -85.0322
4.795 | -85.032
4.7953 | -85.0318
4.7956 | -85.0318
4.7959 | -85.0314
4.7963 | -85.0304
4.7966 | -85.0296
4.7969 | -85.0292
4.7972 | -85.0285
4.7975 | -85.0278
4.7978 | -85.0269
4.7981 | -85.0263
4.7984 | -85.026
4.7987 | -85.0258
4.7991 | -85.0253
4.7994 | -85.0252
4.7997 | -85.0259
4.8 | -85.0263
4.8003 | -85.0257
4.8006 | -85.0251
4.8009 | -85.0251
4.8012 | -85.0247
4.8015 | -85.0255
4.8018 | -85.0258
4.8022 | -85.0243
4.8025 | -85.0221
4.8028 | -85.0215
4.8031 | -85.0219
4.8034 | -85.0212
4.8037 | -85.021
4.804 | -85.0207
4.8043 | -85.0194
4.8046 | -85.0193
4.805 | -85.0188
4.8053 | -85.0185
4.8056 | -85.0173
4.8059 | -85.0152
4.8062 | -85.0146
4.8065 | -85.0157
4.8068 | -85.0163
4.8071 | -85.0155
4.8074 | -85.0147
4.8077 | -85.0144
4.8081 | -85.0141
4.8084 | -85.014
4.8087 | -85.0147
4.809 | -85.015
4.8093 | -85.0145
4.8096 | -85.0141
4.8099 | -85.014
4.8102 | -85.0139
4.8105 | -85.0135
4.8109 | -85.013
4.8112 | -85.0119
4.8115 | -85.0116
4.8118 | -85.0134
4.8121 | -85.0152
4.8124 | -85.0154
4.8127 | -85.0153
4.813 | -85.015
4.8133 | -85.0145
4.8136 | -85.0147
4.814 | -85.0147
4.8143 | -85.0151
4.8146 | -85.0158
4.8149 | -85.0153
4.8152 | -85.0144
4.8155 | -85.0125
4.8158 | -85.0111
4.8161 | -85.0105
4.8164 | -85.0099
4.8168 | -85.0094
4.8171 | -85.0088
4.8174 | -85.0087
4.8177 | -85.0082
4.818 | -85.0082
4.8183 | -85.0075
4.8186 | -85.006
4.8189 | -85.0044
4.8192 | -85.0037
4.8195 | -85.0037
4.8199 | -85.0034
4.8202 | -85.002
4.8205 | -85
4.8208 | -84.9994
4.8211 | -84.9987
4.8214 | -84.9986
4.8217 | -84.9992
4.822 | -84.9997
4.8223 | -84.9998
4.8227 | -84.9991
4.823 | -84.9985
4.8233 | -84.9982
4.8236 | -84.999
4.8239 | -85.0001
4.8242 | -85.0001
4.8245 | -84.999
4.8248 | -84.9972
4.8251 | -84.9967
4.8254 | -84.9966
4.8258 | -84.9959
4.8261 | -84.9944
4.8264 | -84.9934
4.8267 | -84.9929
4.827 | -84.9929
4.8273 | -84.9922
4.8276 | -84.9908
4.8279 | -84.9893
4.8282 | -84.9894
4.8286 | -84.9896
4.8289 | -84.9892
4.8292 | -84.989
4.8295 | -84.9884
4.8298 | -84.9877
4.8301 | -84.9855
4.8304 | -84.9849
4.8307 | -84.9845
4.831 | -84.984
4.8313 | -84.9834
4.8317 | -84.9822
4.832 | -84.9819
4.8323 | -84.9798
4.8326 | -84.9794
4.8329 | -84.9796
4.8332 | -84.9794
4.8335 | -84.9787
4.8338 | -84.9789
4.8341 | -84.9789
4.8345 | -84.9783
4.8348 | -84.9775
4.8351 | -84.9766
4.8354 | -84.9758
4.8357 | -84.9753
4.836 | -84.9747
4.8363 | -84.9732
4.8366 | -84.9729
4.8369 | -84.9726
4.8372 | -84.9723
4.8376 | -84.9725
4.8379 | -84.9724
4.8382 | -84.9724
4.8385 | -84.9718
4.8388 | -84.9715
4.8391 | -84.9708
4.8394 | -84.9692
4.8397 | -84.9678
4.84 | -84.9675
4.8404 | -84.9675
4.8407 | -84.9673
4.841 | -84.9667
4.8413 | -84.9668
4.8416 | -84.9676
4.8419 | -84.9673
4.8422 | -84.9665
4.8425 | -84.9657
4.8428 | -84.9657
4.8431 | -84.966
4.8435 | -84.9664
4.8438 | -84.9664
4.8441 | -84.9661
4.8444 | -84.9653
4.8447 | -84.9643
4.845 | -84.9641
4.8453 | -84.9641
4.8456 | -84.9631
4.8459 | -84.9628
4.8463 | -84.9625
4.8466 | -84.962
4.8469 | -84.9608
4.8472 | -84.9602
4.8475 | -84.9595
4.8478 | -84.9581
4.8481 | -84.9574
4.8484 | -84.9572
4.8487 | -84.9578
4.849 | -84.9595
4.8494 | -84.9602
4.8497 | -84.9603
4.85 | -84.9619
4.8503 | -84.9638
4.8506 | -84.9648
4.8509 | -84.9644
4.8512 | -84.9639
4.8515 | -84.9634
4.8518 | -84.9626
4.8522 | -84.9616
4.8525 | -84.9611
4.8528 | -84.9594
4.8531 | -84.9579
4.8534 | -84.9578
4.8537 | -84.9585
4.854 | -84.958
4.8543 | -84.9571
4.8546 | -84.9562
4.8549 | -84.9544
4.8553 | -84.9531
4.8556 | -84.9523
4.8559 | -84.9518
4.8562 | -84.9516
4.8565 | -84.9509
4.8568 | -84.9506
4.8571 | -84.9504
4.8574 | -84.9494
4.8577 | -84.9485
4.8581 | -84.9478
4.8584 | -84.9473
4.8587 | -84.9469
4.859 | -84.9463
4.8593 | -84.9454
4.8596 | -84.9456
4.8599 | -84.9463
4.8602 | -84.9462
4.8605 | -84.9459
4.8608 | -84.9458
4.8612 | -84.9448
4.8615 | -84.9449
4.8618 | -84.945
4.8621 | -84.9444
4.8624 | -84.944
4.8627 | -84.9434
4.863 | -84.9427
4.8633 | -84.9423
4.8636 | -84.942
4.864 | -84.942
4.8643 | -84.9424
4.8646 | -84.9421
4.8649 | -84.9418
4.8652 | -84.9411
4.8655 | -84.9407
4.8658 | -84.9411
4.8661 | -84.9407
4.8664 | -84.9399
4.8667 | -84.9396
4.8671 | -84.9391
4.8674 | -84.9388
4.8677 | -84.9382
4.868 | -84.9374
4.8683 | -84.9366
4.8686 | -84.9364
4.8689 | -84.9357
4.8692 | -84.935
4.8695 | -84.9349
4.8699 | -84.9345
4.8702 | -84.9343
4.8705 | -84.9342
4.8708 | -84.9338
4.8711 | -84.934
4.8714 | -84.9336
4.8717 | -84.9333
4.872 | -84.9337
4.8723 | -84.933
4.8726 | -84.9331
4.873 | -84.9339
4.8733 | -84.9338
4.8736 | -84.9336
4.8739 | -84.934
4.8742 | -84.9337
4.8745 | -84.9347
4.8748 | -84.9361
4.8751 | -84.9372
4.8754 | -84.9375
4.8758 | -84.9379
4.8761 | -84.9382
4.8764 | -84.938
4.8767 | -84.9379
4.877 | -84.9372
4.8773 | -84.937
4.8776 | -84.9354
4.8779 | -84.934
4.8782 | -84.9331
4.8785 | -84.932
4.8789 | -84.9312
4.8792 | -84.9315
4.8795 | -84.9314
4.8798 | -84.931
4.8801 | -84.9306
4.8804 | -84.9303
4.8807 | -84.9295
4.881 | -84.9287
4.8813 | -84.9282
4.8817 | -84.928
4.882 | -84.9279
4.8823 | -84.9282
4.8826 | -84.9282
4.8829 | -84.9272
4.8832 | -84.9274
4.8835 | -84.9274
4.8838 | -84.9281
4.8841 | -84.929
4.8844 | -84.9292
4.8848 | -84.9289
4.8851 | -84.9286
4.8854 | -84.9285
4.8857 | -84.9282
4.886 | -84.9281
4.8863 | -84.9288
4.8866 | -84.9284
4.8869 | -84.9278
4.8872 | -84.9273
4.8876 | -84.9286
4.8879 | -84.9292
4.8882 | -84.929
4.8885 | -84.9294
4.8888 | -84.9298
4.8891 | -84.9291
4.8894 | -84.9289
4.8897 | -84.9297
4.89 | -84.93
4.8903 | -84.93
4.8907 | -84.93
4.891 | -84.9296
4.8913 | -84.9297
4.8916 | -84.9304
4.8919 | -84.9311
4.8922 | -84.931
4.8925 | -84.9304
4.8928 | -84.9302
4.8931 | -84.9297
4.8935 | -84.9297
4.8938 | -84.9297
4.8941 | -84.9303
4.8944 | -84.9308
4.8947 | -84.9306
4.895 | -84.93
4.8953 | -84.9295
4.8956 | -84.9292
4.8959 | -84.9282
4.8962 | -84.9279
4.8966 | -84.928
4.8969 | -84.929
4.8972 | -84.9301
4.8975 | -84.931
4.8978 | -84.9311
4.8981 | -84.9305
4.8984 | -84.9301
4.8987 | -84.9302
4.899 | -84.9299
4.8994 | -84.93
4.8997 | -84.9301
4.9 | -84.9302
4.9003 | -84.9303
4.9006 | -84.9307
4.9009 | -84.932
4.9012 | -84.9327
4.9015 | -84.9333
4.9018 | -84.9331
4.9021 | -84.9332
4.9025 | -84.9335
4.9028 | -84.9335
4.9031 | -84.9336
4.9034 | -84.9336
4.9037 | -84.9343
4.904 | -84.9351
4.9043 | -84.9355
4.9046 | -84.9354
4.9049 | -84.9351
4.9053 | -84.9352
4.9056 | -84.9346
4.9059 | -84.9341
4.9062 | -84.9341
4.9065 | -84.934
4.9068 | -84.9336
4.9071 | -84.9317
4.9074 | -84.9311
4.9077 | -84.9314
4.9081 | -84.9311
4.9084 | -84.9308
4.9087 | -84.9327
4.909 | -84.9333
4.9093 | -84.9316
4.9096 | -84.9302
4.9099 | -84.9297
4.9102 | -84.9308
4.9105 | -84.9314
4.9108 | -84.9323
4.9112 | -84.9329
4.9115 | -84.933
4.9118 | -84.9338
4.9121 | -84.9346
4.9124 | -84.936
4.9127 | -84.936
4.913 | -84.9361
4.9133 | -84.9357
4.9136 | -84.9357
4.9139 | -84.936
4.9143 | -84.9375
4.9146 | -84.9381
4.9149 | -84.9383
4.9152 | -84.9385
4.9155 | -84.9387
4.9158 | -84.9393
4.9161 | -84.9404
4.9164 | -84.9405
4.9167 | -84.9415
4.9171 | -84.9416
4.9174 | -84.942
4.9177 | -84.9443
4.918 | -84.9459
4.9183 | -84.9462
4.9186 | -84.9474
4.9189 | -84.9476
4.9192 | -84.9477
4.9195 | -84.948
4.9199 | -84.9459
4.9202 | -84.944
4.9205 | -84.9444
4.9208 | -84.9438
4.9211 | -84.9429
4.9214 | -84.943
4.9217 | -84.9432
4.922 | -84.9435
4.9223 | -84.9447
4.9226 | -84.945
4.923 | -84.9443
4.9233 | -84.9439
4.9236 | -84.944
4.9239 | -84.9441
4.9242 | -84.9444
4.9245 | -84.9446
4.9248 | -84.944
4.9251 | -84.9434
4.9254 | -84.9433
4.9257 | -84.9434
4.9261 | -84.9427
4.9264 | -84.9431
4.9267 | -84.9439
4.927 | -84.9438
4.9273 | -84.9441
4.9276 | -84.9451
4.9279 | -84.9446
4.9282 | -84.9447
4.9285 | -84.9451
4.9289 | -84.9451
4.9292 | -84.9448
4.9295 | -84.9443
4.9298 | -84.9441
4.9301 | -84.9451
4.9304 | -84.9469
4.9307 | -84.9479
4.931 | -84.9485
4.9313 | -84.9486
4.9317 | -84.9491
4.932 | -84.9496
4.9323 | -84.9501
4.9326 | -84.9501
4.9329 | -84.9502
4.9332 | -84.95
4.9335 | -84.9501
4.9338 | -84.9506
4.9341 | -84.9509
4.9344 | -84.9512
4.9348 | -84.9517
4.9351 | -84.9509
4.9354 | -84.9499
4.9357 | -84.9496
4.936 | -84.9499
4.9363 | -84.9507
4.9366 | -84.9508
4.9369 | -84.9504
4.9372 | -84.9512
4.9376 | -84.9517
4.9379 | -84.952
4.9382 | -84.9522
4.9385 | -84.9528
4.9388 | -84.9525
4.9391 | -84.9526
4.9394 | -84.9538
4.9397 | -84.9546
4.94 | -84.955
4.9403 | -84.9552
4.9407 | -84.9563
4.941 | -84.9568
4.9413 | -84.957
4.9416 | -84.9571
4.9419 | -84.9578
4.9422 | -84.9582
4.9425 | -84.9592
4.9428 | -84.9597
4.9431 | -84.9603
4.9435 | -84.9606
4.9438 | -84.9603
4.9441 | -84.96
4.9444 | -84.9594
4.9447 | -84.9596
4.945 | -84.9608
4.9453 | -84.961
4.9456 | -84.9617
4.9459 | -84.9622
4.9462 | -84.9626
4.9466 | -84.9631
4.9469 | -84.9631
4.9472 | -84.9637
4.9475 | -84.9643
4.9478 | -84.9645
4.9481 | -84.9655
4.9484 | -84.966
4.9487 | -84.9663
4.949 | -84.9664
4.9494 | -84.9658
4.9497 | -84.9658
4.95 | -84.9663
4.9503 | -84.9673
4.9506 | -84.9682
4.9509 | -84.9683
4.9512 | -84.9673
4.9515 | -84.9674
4.9518 | -84.9676
4.9521 | -84.968
4.9525 | -84.9684
4.9528 | -84.9688
4.9531 | -84.9694
4.9534 | -84.9711
4.9537 | -84.9713
4.954 | -84.9709
4.9543 | -84.9703
4.9546 | -84.9703
4.9549 | -84.9696
4.9553 | -84.9686
4.9556 | -84.9684
4.9559 | -84.9687
4.9562 | -84.9687
4.9565 | -84.9691
4.9568 | -84.9684
4.9571 | -84.9687
4.9574 | -84.9696
4.9577 | -84.9697
4.958 | -84.9702
4.9584 | -84.9706
4.9587 | -84.9699
4.959 | -84.9708
4.9593 | -84.9711
4.9596 | -84.9715
4.9599 | -84.9719
4.9602 | -84.9721
4.9605 | -84.9727
4.9608 | -84.9737
4.9612 | -84.9739
4.9615 | -84.9743
4.9618 | -84.9748
4.9621 | -84.9755
4.9624 | -84.9759
4.9627 | -84.9763
4.963 | -84.9767
4.9633 | -84.9772
4.9636 | -84.9784
4.9639 | -84.9793
4.9643 | -84.9796
4.9646 | -84.9806
4.9649 | -84.9817
4.9652 | -84.9822
4.9655 | -84.9827
4.9658 | -84.9838
4.9661 | -84.9837
4.9664 | -84.9838
4.9667 | -84.9844
4.9671 | -84.9849
4.9674 | -84.9853
4.9677 | -84.9854
4.968 | -84.9863
4.9683 | -84.9868
4.9686 | -84.9874
4.9689 | -84.9879
4.9692 | -84.9885
4.9695 | -84.9893
4.9698 | -84.9888
4.9702 | -84.989
4.9705 | -84.9897
4.9708 | -84.9907
4.9711 | -84.9915
4.9714 | -84.9921
4.9717 | -84.9924
4.972 | -84.9946
4.9723 | -84.9956
4.9726 | -84.9963
4.973 | -84.9964
4.9733 | -84.9964
4.9736 | -84.997
4.9739 | -84.9971
4.9742 | -84.9977
4.9745 | -84.9981
4.9748 | -84.9982
4.9751 | -84.9987
4.9754 | -84.999
4.9757 | -84.999
4.9761 | -84.9989
4.9764 | -85.0003
4.9767 | -85.0012
4.977 | -85.0002
4.9773 | -85.0005
4.9776 | -84.9999
4.9779 | -84.9999
4.9782 | -85.0002
4.9785 | -85.0007
4.9789 | -85.0011
4.9792 | -85.0015
4.9795 | -85.0019
4.9798 | -85.0025
4.9801 | -85.0032
4.9804 | -85.0035
4.9807 | -85.0032
4.981 | -85.0032
4.9813 | -85.0041
4.9816 | -85.0045
4.982 | -85.0051
4.9823 | -85.0059
4.9826 | -85.0067
4.9829 | -85.0074
4.9832 | -85.008
4.9835 | -85.0075
4.9838 | -85.0087
4.9841 | -85.0105
4.9844 | -85.0107
4.9848 | -85.0103
4.9851 | -85.0099
4.9854 | -85.0094
4.9857 | -85.0095
4.986 | -85.0099
4.9863 | -85.0101
4.9866 | -85.0107
4.9869 | -85.0116
4.9872 | -85.0122
4.9875 | -85.0129
4.9879 | -85.0132
4.9882 | -85.0137
4.9885 | -85.0141
4.9888 | -85.0157
4.9891 | -85.0171
4.9894 | -85.0176
4.9897 | -85.0181
4.99 | -85.019
4.9903 | -85.0197
4.9907 | -85.0201
4.991 | -85.0206
4.9913 | -85.0211
4.9916 | -85.0218
4.9919 | -85.0225
4.9922 | -85.0226
4.9925 | -85.0226
4.9928 | -85.0234
4.9931 | -85.0241
4.9934 | -85.0245
4.9938 | -85.0249
4.9941 | -85.0256
4.9944 | -85.0249
4.9947 | -85.0253
4.995 | -85.0261
4.9953 | -85.0266
4.9956 | -85.0262
4.9959 | -85.0269
4.9962 | -85.0272
4.9966 | -85.0275
4.9969 | -85.0279
4.9972 | -85.028
4.9975 | -85.0282
4.9978 | -85.0283
4.9981 | -85.0285
4.9984 | -85.0284
4.9987 | -85.0281
4.999 | -85.0285
4.9993 | -85.0288
4.9997 | -85.0293
5 | -85.0298
5.0003 | -85.0301
5.0006 | -85.0305
5.0009 | -85.0307
5.0012 | -85.031
5.0015 | -85.0314
5.0018 | -85.0353
5.0021 | -85.0412
5.0025 | -85.0426
5.0028 | -85.0427
5.0031 | -85.0435
5.0034 | -85.0438
5.0037 | -85.0431
5.004 | -85.0442
5.0043 | -85.045
5.0046 | -85.0452
5.0049 | -85.0456
5.0052 | -85.0457
5.0056 | -85.0448
5.0059 | -85.0448
5.0062 | -85.0453
5.0065 | -85.0461
5.0068 | -85.0464
5.0071 | -85.046
5.0074 | -85.046
5.0077 | -85.0469
5.008 | -85.0479
5.0084 | -85.0484
5.0087 | -85.0496
5.009 | -85.0503
5.0093 | -85.0507
5.0096 | -85.052
5.0099 | -85.0536
5.0102 | -85.0541
5.0105 | -85.0565
5.0108 | -85.059
5.0111 | -85.0598
5.0115 | -85.0599
5.0118 | -85.0607
5.0121 | -85.0616
5.0124 | -85.0627
5.0127 | -85.0631
5.013 | -85.0625
5.0133 | -85.0628
5.0136 | -85.0642
5.0139 | -85.0647
5.0143 | -85.0647
5.0146 | -85.065
5.0149 | -85.0647
5.0152 | -85.0655
5.0155 | -85.0666
5.0158 | -85.0671
5.0161 | -85.0673
5.0164 | -85.0677
5.0167 | -85.0677
5.017 | -85.0689
5.0174 | -85.0698
5.0177 | -85.0709
5.018 | -85.0707
5.0183 | -85.0705
5.0186 | -85.0715
5.0189 | -85.0722
5.0192 | -85.0723
5.0195 | -85.0727
5.0198 | -85.0732
5.0202 | -85.0731
5.0205 | -85.0735
5.0208 | -85.0738
5.0211 | -85.0739
5.0214 | -85.0747
5.0217 | -85.0764
5.022 | -85.0766
5.0223 | -85.0774
5.0226 | -85.0774
5.0229 | -85.078
5.0233 | -85.0783
5.0236 | -85.0774
5.0239 | -85.0768
5.0242 | -85.0769
5.0245 | -85.0767
5.0248 | -85.0767
5.0251 | -85.0772
5.0254 | -85.0771
5.0257 | -85.077
5.0261 | -85.0772
5.0264 | -85.0776
5.0267 | -85.0776
5.027 | -85.0773
5.0273 | -85.0773
5.0276 | -85.0769
5.0279 | -85.0761
5.0282 | -85.0764
5.0285 | -85.0769
5.0288 | -85.0778
5.0292 | -85.0791
5.0295 | -85.0806
5.0298 | -85.0807
5.0301 | -85.0802
5.0304 | -85.0806
5.0307 | -85.0808
5.031 | -85.081
5.0313 | -85.082
5.0316 | -85.0828
5.032 | -85.0828
5.0323 | -85.083
5.0326 | -85.0834
5.0329 | -85.0841
5.0332 | -85.0854
5.0335 | -85.0876
5.0338 | -85.0888
5.0341 | -85.0892
5.0344 | -85.0895
5.0347 | -85.0895
5.0351 | -85.0892
5.0354 | -85.0888
5.0357 | -85.0884
5.036 | -85.0886
5.0363 | -85.0896
5.0366 | -85.0888
5.0369 | -85.0879
5.0372 | -85.088
5.0375 | -85.0887
5.0379 | -85.0891
5.0382 | -85.09
5.0385 | -85.0903
5.0388 | -85.0906
5.0391 | -85.0908
5.0394 | -85.0914
5.0397 | -85.0923
5.04 | -85.0928
5.0403 | -85.0937
5.0406 | -85.0944
5.041 | -85.0956
5.0413 | -85.0961
5.0416 | -85.0963
5.0419 | -85.0965
5.0422 | -85.097
5.0425 | -85.0975
5.0428 | -85.0977
5.0431 | -85.0975
5.0434 | -85.0972
5.0438 | -85.0974
5.0441 | -85.0983
5.0444 | -85.0978
5.0447 | -85.0982
5.045 | -85.0981
5.0453 | -85.0968
5.0456 | -85.0968
5.0459 | -85.0968
5.0462 | -85.097
5.0465 | -85.0972
5.0469 | -85.0975
5.0472 | -85.0978
5.0475 | -85.0985
5.0478 | -85.099
5.0481 | -85.0996
5.0484 | -85.1001
5.0487 | -85.0993
5.049 | -85.0971
5.0493 | -85.0948
5.0497 | -85.0949
5.05 | -85.0954
5.0503 | -85.0957
5.0506 | -85.0965
5.0509 | -85.0968
5.0512 | -85.0971
5.0515 | -85.0974
5.0518 | -85.0971
5.0521 | -85.0972
5.0525 | -85.0968
5.0528 | -85.0949
5.0531 | -85.0941
5.0534 | -85.0949
5.0537 | -85.0948
5.054 | -85.0948
5.0543 | -85.0951
5.0546 | -85.0949
5.0549 | -85.0946
5.0552 | -85.095
5.0556 | -85.0954
5.0559 | -85.0952
5.0562 | -85.0953
5.0565 | -85.0955
5.0568 | -85.0958
5.0571 | -85.0964
5.0574 | -85.0979
5.0577 | -85.0979
5.058 | -85.0983
5.0583 | -85.0987
5.0587 | -85.0995
5.059 | -85.1005
5.0593 | -85.1015
5.0596 | -85.1029
5.0599 | -85.104
5.0602 | -85.1045
5.0605 | -85.1046
5.0608 | -85.1044
5.0611 | -85.1042
5.0615 | -85.104
5.0618 | -85.1039
5.0621 | -85.105
5.0624 | -85.1055
5.0627 | -85.1058
5.063 | -85.105
5.0633 | -85.104
5.0636 | -85.1043
5.0639 | -85.1046
5.0643 | -85.1047
5.0646 | -85.1054
5.0649 | -85.1059
5.0652 | -85.1064
5.0655 | -85.1067
5.0658 | -85.1069
5.0661 | -85.107
5.0664 | -85.1075
5.0667 | -85.11
5.067 | -85.1105
5.0674 | -85.1109
5.0677 | -85.1119
5.068 | -85.1125
5.0683 | -85.1129
5.0686 | -85.1131
5.0689 | -85.1135
5.0692 | -85.1135
5.0695 | -85.1138
5.0698 | -85.1138
5.0701 | -85.1137
5.0705 | -85.1139
5.0708 | -85.1127
5.0711 | -85.113
5.0714 | -85.1136
5.0717 | -85.1141
5.072 | -85.1143
5.0723 | -85.1133
5.0726 | -85.1118
5.0729 | -85.1115
5.0733 | -85.1129
5.0736 | -85.1137
5.0739 | -85.114
5.0742 | -85.1148
5.0745 | -85.1163
5.0748 | -85.117
5.0751 | -85.1175
5.0754 | -85.1183
5.0757 | -85.1193
5.0761 | -85.1196
5.0764 | -85.1203
5.0767 | -85.1211
5.077 | -85.1212
5.0773 | -85.1216
5.0776 | -85.1224
5.0779 | -85.1236
5.0782 | -85.1241
5.0785 | -85.1255
5.0788 | -85.1283
5.0792 | -85.1308
5.0795 | -85.1315
5.0798 | -85.1321
5.0801 | -85.1323
5.0804 | -85.1329
5.0807 | -85.1327
5.081 | -85.1333
5.0813 | -85.1343
5.0816 | -85.1347
5.082 | -85.1354
5.0823 | -85.1358
5.0826 | -85.1364
5.0829 | -85.1376
5.0832 | -85.1385
5.0835 | -85.1394
5.0838 | -85.1402
5.0841 | -85.14
5.0844 | -85.1408
5.0847 | -85.1412
5.0851 | -85.1417
5.0854 | -85.1423
5.0857 | -85.1425
5.086 | -85.1424
5.0863 | -85.1428
5.0866 | -85.1438
5.0869 | -85.1455
5.0872 | -85.1461
5.0875 | -85.1463
5.0879 | -85.1465
5.0882 | -85.1468
5.0885 | -85.1489
5.0888 | -85.1509
5.0891 | -85.1524
5.0894 | -85.1531
5.0897 | -85.154
5.09 | -85.1542
5.0903 | -85.1538
5.0906 | -85.1551
5.091 | -85.1564
5.0913 | -85.1568
5.0916 | -85.1577
5.0919 | -85.1581
5.0922 | -85.1587
5.0925 | -85.1594
5.0928 | -85.1602
5.0931 | -85.1611
5.0934 | -85.162
5.0938 | -85.1628
5.0941 | -85.1641
5.0944 | -85.1656
5.0947 | -85.1667
5.095 | -85.1668
5.0953 | -85.1672
5.0956 | -85.1672
5.0959 | -85.1684
5.0962 | -85.169
5.0965 | -85.1705
5.0969 | -85.1719
5.0972 | -85.173
5.0975 | -85.1736
5.0978 | -85.1755
5.0981 | -85.177
5.0984 | -85.181
5.0987 | -85.184
5.099 | -85.1866
5.0993 | -85.1892
5.0997 | -85.1902
5.1 | -85.1915
5.1003 | -85.1927
5.1006 | -85.1938
5.1009 | -85.1938
5.1012 | -85.1941
5.1015 | -85.1955
5.1018 | -85.1966
5.1021 | -85.198
5.1024 | -85.1986
5.1028 | -85.1991
5.1031 | -85.1999
5.1034 | -85.2005
5.1037 | -85.2022
5.104 | -85.2037
5.1043 | -85.2035
5.1046 | -85.2043
5.1049 | -85.2056
5.1052 | -85.2062
5.1056 | -85.2071
5.1059 | -85.2078
5.1062 | -85.2091
5.1065 | -85.2104
5.1068 | -85.212
5.1071 | -85.2127
5.1074 | -85.2132
5.1077 | -85.214
5.108 | -85.215
5.1083 | -85.2158
5.1087 | -85.2161
5.109 | -85.2165
5.1093 | -85.2168
5.1096 | -85.2173
5.1099 | -85.2183
5.1102 | -85.219
5.1105 | -85.2197
5.1108 | -85.2214
5.1111 | -85.2263
5.1115 | -85.229
5.1118 | -85.2312
5.1121 | -85.2319
5.1124 | -85.2326
5.1127 | -85.2338
5.113 | -85.2344
5.1133 | -85.2346
5.1136 | -85.235
5.1139 | -85.2362
5.1142 | -85.2384
5.1146 | -85.2403
5.1149 | -85.2415
5.1152 | -85.2424
5.1155 | -85.2426
5.1158 | -85.2443
5.1161 | -85.2453
5.1164 | -85.2465
5.1167 | -85.2477
5.117 | -85.2491
5.1174 | -85.2507
5.1177 | -85.252
5.118 | -85.2529
5.1183 | -85.2536
5.1186 | -85.2545
5.1189 | -85.2554
5.1192 | -85.2565
5.1195 | -85.2579
5.1198 | -85.259
5.1201 | -85.2611
5.1205 | -85.2635
5.1208 | -85.2648
5.1211 | -85.2664
5.1214 | -85.2684
5.1217 | -85.2693
5.122 | -85.2703
5.1223 | -85.2719
5.1226 | -85.2737
5.1229 | -85.2747
5.1233 | -85.2756
5.1236 | -85.2762
5.1239 | -85.2776
5.1242 | -85.2792
5.1245 | -85.2808
5.1248 | -85.2822
5.1251 | -85.2836
5.1254 | -85.2849
5.1257 | -85.2862
5.126 | -85.2875
5.1264 | -85.2886
5.1267 | -85.2894
5.127 | -85.2909
5.1273 | -85.2931
5.1276 | -85.295
5.1279 | -85.2958
5.1282 | -85.2961
5.1285 | -85.2974
5.1288 | -85.299
5.1292 | -85.3002
5.1295 | -85.3013
5.1298 | -85.3023
5.1301 | -85.3032
5.1304 | -85.3036
5.1307 | -85.3053
5.131 | -85.3071
5.1313 | -85.3086
5.1316 | -85.3095
5.1319 | -85.3104
5.1323 | -85.3117
5.1326 | -85.3132
5.1329 | -85.3148
5.1332 | -85.3155
5.1335 | -85.3164
5.1338 | -85.3172
5.1341 | -85.3181
5.1344 | -85.319
5.1347 | -85.3207
5.1351 | -85.3222
5.1354 | -85.3239
5.1357 | -85.3253
5.136 | -85.3265
5.1363 | -85.3287
5.1366 | -85.3314
5.1369 | -85.3317
5.1372 | -85.332
5.1375 | -85.3331
5.1378 | -85.3345
5.1382 | -85.3352
5.1385 | -85.3359
5.1388 | -85.3376
5.1391 | -85.3391
5.1394 | -85.3403
5.1397 | -85.3418
5.14 | -85.3433
5.1403 | -85.3446
5.1406 | -85.3459
5.141 | -85.3472
5.1413 | -85.3487
5.1416 | -85.3507
5.1419 | -85.3525
5.1422 | -85.3549
5.1425 | -85.3576
5.1428 | -85.36
5.1431 | -85.3616
5.1434 | -85.3631
5.1437 | -85.3643
5.1441 | -85.3654
5.1444 | -85.3672
5.1447 | -85.3687
5.145 | -85.3689
5.1453 | -85.3696
5.1456 | -85.3719
5.1459 | -85.3725
5.1462 | -85.373
5.1465 | -85.3748
5.1469 | -85.3763
5.1472 | -85.3776
5.1475 | -85.3789
5.1478 | -85.3818
5.1481 | -85.3853
5.1484 | -85.3871
5.1487 | -85.388
5.149 | -85.3893
5.1493 | -85.3901
5.1496 | -85.3915
5.15 | -85.3942
5.1503 | -85.396
5.1506 | -85.3977
5.1509 | -85.3998
5.1512 | -85.4033
5.1515 | -85.407
5.1518 | -85.4088
5.1521 | -85.4105
5.1524 | -85.4115
5.1528 | -85.4131
5.1531 | -85.4148
5.1534 | -85.4164
5.1537 | -85.4176
5.154 | -85.4186
5.1543 | -85.4202
5.1546 | -85.4215
5.1549 | -85.423
5.1552 | -85.4245
5.1555 | -85.4268
5.1559 | -85.4287
5.1562 | -85.4295
5.1565 | -85.4311
5.1568 | -85.4328
5.1571 | -85.4344
5.1574 | -85.4362
5.1577 | -85.4374
5.158 | -85.4391
5.1583 | -85.4403
5.1587 | -85.4417
5.159 | -85.4432
5.1593 | -85.4446
5.1596 | -85.4462
5.1599 | -85.4482
5.1602 | -85.4499
5.1605 | -85.4514
5.1608 | -85.4525
5.1611 | -85.4541
5.1614 | -85.4558
5.1618 | -85.4578
5.1621 | -85.4613
5.1624 | -85.4633
5.1627 | -85.4649
5.163 | -85.4668
5.1633 | -85.4685
5.1636 | -85.469
5.1639 | -85.4703
5.1642 | -85.4717
5.1646 | -85.4728
5.1649 | -85.4751
5.1652 | -85.4773
5.1655 | -85.4797
5.1658 | -85.4813
5.1661 | -85.4827
5.1664 | -85.4843
5.1667 | -85.4856
5.167 | -85.4875
5.1673 | -85.4906
5.1677 | -85.4926
5.168 | -85.4945
5.1683 | -85.4954
5.1686 | -85.4959
5.1689 | -85.4977
5.1692 | -85.4987
5.1695 | -85.5009
5.1698 | -85.5027
5.1701 | -85.5047
5.1705 | -85.5063
5.1708 | -85.5079
5.1711 | -85.5091
5.1714 | -85.5105
5.1717 | -85.5127
5.172 | -85.5155
5.1723 | -85.517
5.1726 | -85.5186
5.1729 | -85.5206
5.1732 | -85.5227
5.1736 | -85.5239
5.1739 | -85.5252
5.1742 | -85.5269
5.1745 | -85.5277
5.1748 | -85.5301
5.1751 | -85.5328
5.1754 | -85.5341
5.1757 | -85.5356
5.176 | -85.5373
5.1764 | -85.5388
5.1767 | -85.5406
5.177 | -85.5419
5.1773 | -85.5431
5.1776 | -85.5464
5.1779 | -85.5483
5.1782 | -85.5502
5.1785 | -85.5522
5.1788 | -85.5542
5.1791 | -85.5564
5.1795 | -85.5585
5.1798 | -85.5605
5.1801 | -85.564
5.1804 | -85.5664
5.1807 | -85.5682
5.181 | -85.57
5.1813 | -85.5716
5.1816 | -85.5731
5.1819 | -85.5745
5.1823 | -85.5762
5.1826 | -85.5777
5.1829 | -85.5788
5.1832 | -85.5809
5.1835 | -85.5826
5.1838 | -85.5831
5.1841 | -85.5842
5.1844 | -85.586
5.1847 | -85.5875
5.185 | -85.5886
5.1854 | -85.5897
5.1857 | -85.5915
5.186 | -85.593
5.1863 | -85.5949
5.1866 | -85.5967
5.1869 | -85.5978
5.1872 | -85.5995
5.1875 | -85.6017
5.1878 | -85.6045
5.1882 | -85.6069
5.1885 | -85.608
5.1888 | -85.6095
5.1891 | -85.6105
5.1894 | -85.6115
5.1897 | -85.6134
5.19 | -85.6155
5.1903 | -85.6176
5.1906 | -85.6188
5.1909 | -85.6196
5.1913 | -85.6203
5.1916 | -85.6221
5.1919 | -85.6239
5.1922 | -85.6249
5.1925 | -85.6254
5.1928 | -85.6261
5.1931 | -85.6272
5.1934 | -85.6289
5.1937 | -85.6308
5.1941 | -85.633
5.1944 | -85.6346
5.1947 | -85.6359
5.195 | -85.6373
5.1953 | -85.6394
5.1956 | -85.6418
5.1959 | -85.6435
5.1962 | -85.645
5.1965 | -85.6471
5.1969 | -85.6501
5.1972 | -85.6537
5.1975 | -85.6562
5.1978 | -85.6577
5.1981 | -85.6588
5.1984 | -85.6602
5.1987 | -85.6621
5.199 | -85.6637
5.1993 | -85.6648
5.1996 | -85.666
5.2 | -85.6684
5.2003 | -85.67
5.2006 | -85.6712
5.2009 | -85.6722
5.2012 | -85.6741
5.2015 | -85.6758
5.2018 | -85.6785
5.2021 | -85.6807
5.2024 | -85.6816
5.2027 | -85.6827
5.2031 | -85.6843
5.2034 | -85.6851
5.2037 | -85.6856
5.204 | -85.6877
5.2043 | -85.6898
5.2046 | -85.6911
5.2049 | -85.6923
5.2052 | -85.693
5.2055 | -85.6914
5.2059 | -85.6933
5.2062 | -85.6944
5.2065 | -85.6959
5.2068 | -85.6972
5.2071 | -85.6991
5.2074 | -85.7007
5.2077 | -85.7023
5.208 | -85.7039
5.2083 | -85.7057
5.2086 | -85.7081
5.209 | -85.711
5.2093 | -85.7123
5.2096 | -85.7133
5.2099 | -85.7148
5.2102 | -85.7158
5.2105 | -85.7173
5.2108 | -85.7192
5.2111 | -85.7204
5.2114 | -85.7217
5.2118 | -85.7234
5.2121 | -85.7249
5.2124 | -85.7264
5.2127 | -85.728
5.213 | -85.7292
5.2133 | -85.7308
5.2136 | -85.7322
5.2139 | -85.7334
5.2142 | -85.7352
5.2146 | -85.736
5.2149 | -85.738
5.2152 | -85.7397
5.2155 | -85.7414
5.2158 | -85.7432
5.2161 | -85.7448
5.2164 | -85.7467
5.2167 | -85.7483
5.217 | -85.7491
5.2173 | -85.7505
5.2177 | -85.7524
5.218 | -85.7542
5.2183 | -85.7562
5.2186 | -85.758
5.2189 | -85.7582
5.2192 | -85.7599
5.2195 | -85.7621
5.2198 | -85.7632
5.2201 | -85.7645
5.2205 | -85.766
5.2208 | -85.7673
5.2211 | -85.7685
5.2214 | -85.7697
5.2217 | -85.7713
5.222 | -85.7729
5.2223 | -85.7743
5.2226 | -85.7762
5.2229 | -85.778
5.2232 | -85.7796
5.2236 | -85.7813
5.2239 | -85.7828
5.2242 | -85.7842
5.2245 | -85.7853
5.2248 | -85.7869
5.2251 | -85.7882
5.2254 | -85.7889
5.2257 | -85.7905
5.226 | -85.792
5.2264 | -85.7934
5.2267 | -85.7947
5.227 | -85.7964
5.2273 | -85.798
5.2276 | -85.7994
5.2279 | -85.8014
5.2282 | -85.8028
5.2285 | -85.8053
5.2288 | -85.8077
5.2291 | -85.8099
5.2295 | -85.8119
5.2298 | -85.8138
5.2301 | -85.8145
5.2304 | -85.8157
5.2307 | -85.8154
5.231 | -85.8165
5.2313 | -85.8181
5.2316 | -85.8203
5.2319 | -85.8224
5.2323 | -85.823
5.2326 | -85.8246
5.2329 | -85.8264
5.2332 | -85.8278
5.2335 | -85.8291
5.2338 | -85.831
5.2341 | -85.8331
5.2344 | -85.8343
5.2347 | -85.8355
5.235 | -85.8367
5.2354 | -85.8377
5.2357 | -85.8389
5.236 | -85.8399
5.2363 | -85.8417
5.2366 | -85.8431
5.2369 | -85.8441
5.2372 | -85.8454
5.2375 | -85.8476
5.2378 | -85.8483
5.2382 | -85.8496
5.2385 | -85.8519
5.2388 | -85.8536
5.2391 | -85.855
5.2394 | -85.8563
5.2397 | -85.8577
5.24 | -85.8588
5.2403 | -85.8604
5.2406 | -85.862
5.2409 | -85.8638
5.2413 | -85.8652
5.2416 | -85.8667
5.2419 | -85.8679
5.2422 | -85.8691
5.2425 | -85.8702
5.2428 | -85.8724
5.2431 | -85.8744
5.2434 | -85.8755
5.2437 | -85.8768
5.2441 | -85.8781
5.2444 | -85.8789
5.2447 | -85.8799
5.245 | -85.8812
5.2453 | -85.8833
5.2456 | -85.885
5.2459 | -85.8863
5.2462 | -85.8875
5.2465 | -85.8888
5.2468 | -85.89
5.2472 | -85.8916
5.2475 | -85.8931
5.2478 | -85.8939
5.2481 | -85.8961
5.2484 | -85.8979
5.2487 | -85.899
5.249 | -85.8993
5.2493 | -85.8999
5.2496 | -85.9008
5.25 | -85.9024
5.2503 | -85.902
5.2506 | -85.904
5.2509 | -85.9048
5.2512 | -85.9053
5.2515 | -85.9068
5.2518 | -85.9073
5.2521 | -85.9087
5.2524 | -85.9101
5.2527 | -85.9108
5.2531 | -85.9117
5.2534 | -85.9128
5.2537 | -85.9139
5.254 | -85.9151
5.2543 | -85.9161
5.2546 | -85.9175
5.2549 | -85.9183
5.2552 | -85.9187
5.2555 | -85.9199
5.2559 | -85.9205
5.2562 | -85.9216
5.2565 | -85.9228
5.2568 | -85.9239
5.2571 | -85.9251
5.2574 | -85.9266
5.2577 | -85.9279
5.258 | -85.9279
5.2583 | -85.929
5.2586 | -85.9308
5.259 | -85.9324
5.2593 | -85.9338
5.2596 | -85.9362
5.2599 | -85.9394
5.2602 | -85.9431
5.2605 | -85.9452
5.2608 | -85.9467
5.2611 | -85.9484
5.2614 | -85.9496
5.2618 | -85.9503
5.2621 | -85.9514
5.2624 | -85.955
5.2627 | -85.9585
5.263 | -85.9607
5.2633 | -85.9621
5.2636 | -85.9627
5.2639 | -85.9638
5.2642 | -85.9652
5.2645 | -85.9668
5.2649 | -85.9691
5.2652 | -85.9702
5.2655 | -85.9722
5.2658 | -85.9743
5.2661 | -85.9756
5.2664 | -85.9766
5.2667 | -85.9781
5.267 | -85.9786
5.2673 | -85.9795
5.2677 | -85.9805
5.268 | -85.9806
5.2683 | -85.9807
5.2686 | -85.9817
5.2689 | -85.9824
5.2692 | -85.9823
5.2695 | -85.982
5.2698 | -85.9823
5.2701 | -85.9835
5.2704 | -85.9835
5.2708 | -85.9838
5.2711 | -85.9844
5.2714 | -85.9854
5.2717 | -85.9869
5.272 | -85.9881
5.2723 | -85.9887
5.2726 | -85.9897
5.2729 | -85.9908
5.2732 | -85.9918
5.2736 | -85.9925
5.2739 | -85.9959
5.2742 | -85.9992
5.2745 | -86.0001
5.2748 | -86.002
5.2751 | -86.0042
5.2754 | -86.0046
5.2757 | -86.0051
5.276 | -86.0062
5.2763 | -86.0071
5.2767 | -86.0082
5.277 | -86.0097
5.2773 | -86.0107
5.2776 | -86.012
5.2779 | -86.013
5.2782 | -86.0136
5.2785 | -86.0143
5.2788 | -86.0156
5.2791 | -86.0171
5.2795 | -86.0187
5.2798 | -86.0199
5.2801 | -86.0198
5.2804 | -86.0203
5.2807 | -86.0208
5.281 | -86.022
5.2813 | -86.0238
5.2816 | -86.0238
5.2819 | -86.024
5.2822 | -86.0252
5.2826 | -86.0264
5.2829 | -86.0271
5.2832 | -86.0277
5.2835 | -86.0288
5.2838 | -86.0302
5.2841 | -86.0317
5.2844 | -86.033
5.2847 | -86.0337
5.285 | -86.0342
5.2854 | -86.0344
5.2857 | -86.0344
5.286 | -86.0343
5.2863 | -86.0352
5.2866 | -86.0366
5.2869 | -86.0364
5.2872 | -86.0362
5.2875 | -86.0369
5.2878 | -86.037
5.2881 | -86.037
5.2885 | -86.0379
5.2888 | -86.0385
5.2891 | -86.0394
5.2894 | -86.0409
5.2897 | -86.0418
5.29 | -86.0439
5.2903 | -86.0462
5.2906 | -86.047
5.2909 | -86.0477
5.2913 | -86.0483
5.2916 | -86.0493
5.2919 | -86.0494
5.2922 | -86.0504
5.2925 | -86.0505
5.2928 | -86.0507
5.2931 | -86.0513
5.2934 | -86.0519
5.2937 | -86.0529
5.294 | -86.0546
5.2944 | -86.0561
5.2947 | -86.057
5.295 | -86.0569
5.2953 | -86.0565
5.2956 | -86.0567
5.2959 | -86.0576
5.2962 | -86.0576
5.2965 | -86.0569
5.2968 | -86.0558
5.2972 | -86.0549
5.2975 | -86.0552
5.2978 | -86.0556
5.2981 | -86.0553
5.2984 | -86.0551
5.2987 | -86.0557
5.299 | -86.0555
5.2993 | -86.0556
5.2996 | -86.0566
5.2999 | -86.0583
5.3003 | -86.0593
5.3006 | -86.0592
5.3009 | -86.0607
5.3012 | -86.061
5.3015 | -86.0609
5.3018 | -86.0615
5.3021 | -86.062
5.3024 | -86.0628
5.3027 | -86.0639
5.3031 | -86.0642
5.3034 | -86.0649
5.3037 | -86.066
5.304 | -86.0665
5.3043 | -86.0674
5.3046 | -86.0674
5.3049 | -86.0669
5.3052 | -86.0686
5.3055 | -86.0705
5.3058 | -86.071
5.3062 | -86.071
5.3065 | -86.0713
5.3068 | -86.0715
5.3071 | -86.0736
5.3074 | -86.0759
5.3077 | -86.0766
5.308 | -86.0765
5.3083 | -86.0767
5.3086 | -86.0772
5.309 | -86.0777
5.3093 | -86.0779
5.3096 | -86.0785
5.3099 | -86.08
5.3102 | -86.082
5.3105 | -86.0839
5.3108 | -86.0852
5.3111 | -86.0854
5.3114 | -86.0869
5.3117 | -86.0879
5.3121 | -86.0884
5.3124 | -86.089
5.3127 | -86.0893
5.313 | -86.0896
5.3133 | -86.0903
5.3136 | -86.0921
5.3139 | -86.0924
5.3142 | -86.0922
5.3145 | -86.0925
5.3149 | -86.0929
5.3152 | -86.0938
5.3155 | -86.0952
5.3158 | -86.0949
5.3161 | -86.0947
5.3164 | -86.0949
5.3167 | -86.0952
5.317 | -86.0954
5.3173 | -86.0942
5.3176 | -86.0935
5.318 | -86.0942
5.3183 | -86.0949
5.3186 | -86.0954
5.3189 | -86.0955
5.3192 | -86.0957
5.3195 | -86.0957
5.3198 | -86.096
5.3201 | -86.096
5.3204 | -86.0962
5.3208 | -86.0965
5.3211 | -86.097
5.3214 | -86.0977
5.3217 | -86.0982
5.322 | -86.0987
5.3223 | -86.0993
5.3226 | -86.0995
5.3229 | -86.099
5.3232 | -86.0987
5.3235 | -86.0995
5.3239 | -86.1
5.3242 | -86.0997
5.3245 | -86.1
5.3248 | -86.1001
5.3251 | -86.099
5.3254 | -86.0983
5.3257 | -86.0985
5.326 | -86.0985
5.3263 | -86.0982
5.3267 | -86.0989
5.327 | -86.0999
5.3273 | -86.1004
5.3276 | -86.1007
5.3279 | -86.101
5.3282 | -86.1012
5.3285 | -86.1006
5.3288 | -86.1004
5.3291 | -86.1006
5.3294 | -86.101
5.3298 | -86.1015
5.3301 | -86.1011
5.3304 | -86.1006
5.3307 | -86.1
5.331 | -86.0992
5.3313 | -86.0997
5.3316 | -86.0996
5.3319 | -86.1006
5.3322 | -86.1006
5.3326 | -86.0999
5.3329 | -86.0996
5.3332 | -86.1004
5.3335 | -86.1014
5.3338 | -86.1021
5.3341 | -86.1034
5.3344 | -86.1041
5.3347 | -86.1042
5.335 | -86.1061
5.3354 | -86.106
5.3357 | -86.1074
5.336 | -86.1082
5.3363 | -86.1087
5.3366 | -86.1087
5.3369 | -86.1083
5.3372 | -86.1083
5.3375 | -86.1084
5.3378 | -86.1087
5.3381 | -86.1084
5.3385 | -86.1079
5.3388 | -86.1084
5.3391 | -86.109
5.3394 | -86.1089
5.3397 | -86.1078
5.34 | -86.1069
5.3403 | -86.1071
5.3406 | -86.1071
5.3409 | -86.1072
5.3412 | -86.1068
5.3416 | -86.1062
5.3419 | -86.1055
5.3422 | -86.1072
5.3425 | -86.1085
5.3428 | -86.1087
5.3431 | -86.1091
5.3434 | -86.1101
5.3437 | -86.1106
5.344 | -86.11
5.3444 | -86.1105
5.3447 | -86.111
5.345 | -86.1105
5.3453 | -86.1101
5.3456 | -86.1097
5.3459 | -86.1096
5.3462 | -86.1095
5.3465 | -86.1098
5.3468 | -86.11
5.3472 | -86.1103
5.3475 | -86.1101
5.3478 | -86.1092
5.3481 | -86.1091
5.3484 | -86.1091
5.3487 | -86.1089
5.349 | -86.1082
5.3493 | -86.1079
5.3496 | -86.1078
5.3499 | -86.1083
5.3503 | -86.1081
5.3506 | -86.1077
5.3509 | -86.1081
5.3512 | -86.1079
5.3515 | -86.1077
5.3518 | -86.1085
5.3521 | -86.1088
5.3524 | -86.1086
5.3527 | -86.1089
5.353 | -86.1097
5.3534 | -86.1097
5.3537 | -86.1106
5.354 | -86.1106
5.3543 | -86.1107
5.3546 | -86.1111
5.3549 | -86.1118
5.3552 | -86.1117
5.3555 | -86.1119
5.3558 | -86.1117
5.3562 | -86.1114
5.3565 | -86.1114
5.3568 | -86.1111
5.3571 | -86.1107
5.3574 | -86.1108
5.3577 | -86.1109
5.358 | -86.1112
5.3583 | -86.1119
5.3586 | -86.1117
5.359 | -86.1119
5.3593 | -86.112
5.3596 | -86.1118
5.3599 | -86.1122
5.3602 | -86.1113
5.3605 | -86.1114
5.3608 | -86.1116
5.3611 | -86.1119
5.3614 | -86.1125
5.3617 | -86.1117
5.3621 | -86.1116
5.3624 | -86.1115
5.3627 | -86.111
5.363 | -86.1111
5.3633 | -86.1113
5.3636 | -86.1124
5.3639 | -86.1125
5.3642 | -86.1123
5.3645 | -86.1125
5.3649 | -86.1126
5.3652 | -86.1131
5.3655 | -86.114
5.3658 | -86.114
5.3661 | -86.1143
5.3664 | -86.1144
5.3667 | -86.1148
5.367 | -86.1155
5.3673 | -86.1164
5.3676 | -86.1168
5.368 | -86.117
5.3683 | -86.1176
5.3686 | -86.1176
5.3689 | -86.1174
5.3692 | -86.1175
5.3695 | -86.1179
5.3698 | -86.1182
5.3701 | -86.1182
5.3704 | -86.1183
5.3708 | -86.1181
5.3711 | -86.1182
5.3714 | -86.1182
5.3717 | -86.1187
5.372 | -86.1195
5.3723 | -86.1194
5.3726 | -86.1196
5.3729 | -86.1193
5.3732 | -86.12
5.3735 | -86.1204
5.3739 | -86.1203
5.3742 | -86.1205
5.3745 | -86.1207
5.3748 | -86.121
5.3751 | -86.121
5.3754 | -86.121
5.3757 | -86.1218
5.376 | -86.1217
5.3763 | -86.1218
5.3767 | -86.122
5.377 | -86.1219
5.3773 | -86.1222
5.3776 | -86.1217
5.3779 | -86.1207
5.3782 | -86.1206
5.3785 | -86.1207
5.3788 | -86.1204
5.3791 | -86.1212
5.3794 | -86.1219
5.3798 | -86.122
5.3801 | -86.1216
5.3804 | -86.1208
5.3807 | -86.1193
5.381 | -86.1177
5.3813 | -86.1166
5.3816 | -86.1161
5.3819 | -86.116
5.3822 | -86.1162
5.3826 | -86.1164
5.3829 | -86.1163
5.3832 | -86.1159
5.3835 | -86.1156
5.3838 | -86.1159
5.3841 | -86.1162
5.3844 | -86.116
5.3847 | -86.1149
5.385 | -86.1158
5.3853 | -86.1167
5.3857 | -86.1163
5.386 | -86.1166
5.3863 | -86.1168
5.3866 | -86.1158
5.3869 | -86.1148
5.3872 | -86.1152
5.3875 | -86.1155
5.3878 | -86.1155
5.3881 | -86.1158
5.3885 | -86.1159
5.3888 | -86.1163
5.3891 | -86.1162
5.3894 | -86.1166
5.3897 | -86.1179
5.39 | -86.1169
5.3903 | -86.1169
5.3906 | -86.1169
5.3909 | -86.1165
5.3912 | -86.1163
5.3916 | -86.1165
5.3919 | -86.1171
5.3922 | -86.1169
5.3925 | -86.1169
5.3928 | -86.1171
5.3931 | -86.1172
5.3934 | -86.1172
5.3937 | -86.1166
5.394 | -86.1166
5.3944 | -86.1166
5.3947 | -86.1167
5.395 | -86.1163
5.3953 | -86.1158
5.3956 | -86.1151
5.3959 | -86.1144
5.3962 | -86.1141
5.3965 | -86.1142
5.3968 | -86.1138
5.3971 | -86.1141
5.3975 | -86.114
5.3978 | -86.1139
5.3981 | -86.1148
5.3984 | -86.116
5.3987 | -86.116
5.399 | -86.1169
5.3993 | -86.1181
5.3996 | -86.1181
5.3999 | -86.1181
5.4003 | -86.1181
5.4006 | -86.118
5.4009 | -86.1176
5.4012 | -86.1173
5.4015 | -86.1173
5.4018 | -86.1174
5.4021 | -86.1173
5.4024 | -86.1165
5.4027 | -86.1168
5.403 | -86.117
5.4034 | -86.1168
5.4037 | -86.1164
5.404 | -86.1163
5.4043 | -86.116
5.4046 | -86.1161
5.4049 | -86.1158
5.4052 | -86.1156
5.4055 | -86.1155
5.4058 | -86.1159
5.4062 | -86.1166
5.4065 | -86.1168
5.4068 | -86.1169
5.4071 | -86.117
5.4074 | -86.1179
5.4077 | -86.1176
5.408 | -86.119
5.4083 | -86.1205
5.4086 | -86.1202
5.4089 | -86.1203
5.4093 | -86.1207
5.4096 | -86.1206
5.4099 | -86.1205
5.4102 | -86.1201
5.4105 | -86.1201
5.4108 | -86.1202
5.4111 | -86.1201
5.4114 | -86.1198
5.4117 | -86.1195
5.4121 | -86.1207
5.4124 | -86.1228
5.4127 | -86.1253
5.413 | -86.1249
5.4133 | -86.1244
5.4136 | -86.1241
5.4139 | -86.1239
5.4142 | -86.1238
5.4145 | -86.1234
5.4148 | -86.1229
5.4152 | -86.1231
5.4155 | -86.1201
5.4158 | -86.1193
5.4161 | -86.1212
5.4164 | -86.1222
5.4167 | -86.1222
5.417 | -86.1222
5.4173 | -86.1231
5.4176 | -86.1231
5.418 | -86.123
5.4183 | -86.1224
5.4186 | -86.1221
5.4189 | -86.1219
5.4192 | -86.1216
5.4195 | -86.1222
5.4198 | -86.1221
5.4201 | -86.1222
5.4204 | -86.1219
5.4207 | -86.1219
5.4211 | -86.1232
5.4214 | -86.123
5.4217 | -86.123
5.422 | -86.1224
5.4223 | -86.1223
5.4226 | -86.1221
5.4229 | -86.122
5.4232 | -86.122
5.4235 | -86.122
5.4239 | -86.1223
5.4242 | -86.1217
5.4245 | -86.1214
5.4248 | -86.1219
5.4251 | -86.1224
5.4254 | -86.1228
5.4257 | -86.1233
5.426 | -86.1233
5.4263 | -86.123
5.4266 | -86.1227
5.427 | -86.123
5.4273 | -86.1243
5.4276 | -86.1257
5.4279 | -86.1255
5.4282 | -86.1255
5.4285 | -86.1251
5.4288 | -86.1246
5.4291 | -86.1236
5.4294 | -86.1233
5.4298 | -86.1229
5.4301 | -86.1228
5.4304 | -86.1228
5.4307 | -86.1231
5.431 | -86.1231
5.4313 | -86.1227
5.4316 | -86.1225
5.4319 | -86.1222
5.4322 | -86.1221
5.4325 | -86.1211
5.4329 | -86.1195
5.4332 | -86.1194
5.4335 | -86.1184
5.4338 | -86.1172
5.4341 | -86.1167
5.4344 | -86.1168
5.4347 | -86.1174
5.435 | -86.1172
5.4353 | -86.1178
5.4357 | -86.1189
5.436 | -86.1187
5.4363 | -86.1185
5.4366 | -86.1186
5.4369 | -86.119
5.4372 | -86.121
5.4375 | -86.1228
5.4378 | -86.1236
5.4381 | -86.1222
5.4384 | -86.1205
5.4388 | -86.1205
5.4391 | -86.1205
5.4394 | -86.1204
5.4397 | -86.1203
5.44 | -86.1201
5.4403 | -86.1196
5.4406 | -86.1186
5.4409 | -86.1185
5.4412 | -86.1166
5.4416 | -86.1158
5.4419 | -86.116
5.4422 | -86.1161
5.4425 | -86.1167
5.4428 | -86.1163
5.4431 | -86.1155
5.4434 | -86.1145
5.4437 | -86.1136
5.444 | -86.1135
5.4443 | -86.1132
5.4447 | -86.1126
5.445 | -86.1123
5.4453 | -86.1125
5.4456 | -86.113
5.4459 | -86.1125
5.4462 | -86.1126
5.4465 | -86.1128
5.4468 | -86.1127
5.4471 | -86.1127
5.4475 | -86.1129
5.4478 | -86.112
5.4481 | -86.1121
5.4484 | -86.1117
5.4487 | -86.1111
5.449 | -86.1109
5.4493 | -86.1104
5.4496 | -86.1095
5.4499 | -86.1104
5.4502 | -86.1103
5.4506 | -86.1099
5.4509 | -86.1098
5.4512 | -86.1093
5.4515 | -86.1088
5.4518 | -86.1094
5.4521 | -86.1095
5.4524 | -86.1096
5.4527 | -86.1096
5.453 | -86.1093
5.4534 | -86.1098
5.4537 | -86.1103
5.454 | -86.1109
5.4543 | -86.1109
5.4546 | -86.1107
5.4549 | -86.1103
5.4552 | -86.1097
5.4555 | -86.1095
5.4558 | -86.1095
5.4561 | -86.1111
5.4565 | -86.1115
5.4568 | -86.1125
5.4571 | -86.1125
5.4574 | -86.1118
5.4577 | -86.1113
5.458 | -86.1112
5.4583 | -86.1115
5.4586 | -86.1116
5.4589 | -86.1115
5.4593 | -86.1112
5.4596 | -86.1109
5.4599 | -86.1108
5.4602 | -86.1109
5.4605 | -86.1106
5.4608 | -86.1103
5.4611 | -86.1104
5.4614 | -86.1112
5.4617 | -86.1117
5.462 | -86.1131
5.4624 | -86.1134
5.4627 | -86.113
5.463 | -86.1129
5.4633 | -86.1126
5.4636 | -86.1119
5.4639 | -86.1113
5.4642 | -86.1115
5.4645 | -86.1118
5.4648 | -86.1116
5.4652 | -86.1122
5.4655 | -86.112
5.4658 | -86.1118
5.4661 | -86.112
5.4664 | -86.1123
5.4667 | -86.1124
5.467 | -86.1118
5.4673 | -86.1112
5.4676 | -86.1113
5.4679 | -86.1123
5.4683 | -86.1126
5.4686 | -86.1122
5.4689 | -86.1121
5.4692 | -86.1122
5.4695 | -86.1119
5.4698 | -86.1119
5.4701 | -86.1118
5.4704 | -86.1121
5.4707 | -86.1121
5.4711 | -86.1118
5.4714 | -86.1127
5.4717 | -86.1127
5.472 | -86.1131
5.4723 | -86.1135
5.4726 | -86.1137
5.4729 | -86.1137
5.4732 | -86.1144
5.4735 | -86.1146
5.4738 | -86.1153
5.4742 | -86.1163
5.4745 | -86.1165
5.4748 | -86.1167
5.4751 | -86.1171
5.4754 | -86.1172
5.4757 | -86.1174
5.476 | -86.118
5.4763 | -86.1189
5.4766 | -86.1197
5.477 | -86.1206
5.4773 | -86.1212
5.4776 | -86.1212
5.4779 | -86.1216
5.4782 | -86.122
5.4785 | -86.1225
5.4788 | -86.1223
5.4791 | -86.1223
5.4794 | -86.1225
5.4798 | -86.1228
5.4801 | -86.1231
5.4804 | -86.1232
5.4807 | -86.1236
5.481 | -86.1246
5.4813 | -86.1253
5.4816 | -86.1256
5.4819 | -86.1253
5.4822 | -86.1255
5.4825 | -86.1259
5.4829 | -86.1252
5.4832 | -86.1252
5.4835 | -86.1269
5.4838 | -86.1277
5.4841 | -86.1278
5.4844 | -86.127
5.4847 | -86.1265
5.485 | -86.1268
5.4853 | -86.1268
5.4856 | -86.1265
5.486 | -86.1265
5.4863 | -86.1276
5.4866 | -86.1281
5.4869 | -86.1284
5.4872 | -86.1289
5.4875 | -86.1291
5.4878 | -86.1291
5.4881 | -86.1295
5.4884 | -86.1299
5.4888 | -86.1299
5.4891 | -86.1297
5.4894 | -86.13
5.4897 | -86.1306
5.49 | -86.1314
5.4903 | -86.1317
5.4906 | -86.1321
5.4909 | -86.1325
5.4912 | -86.1327
5.4916 | -86.1334
5.4919 | -86.1343
5.4922 | -86.1355
5.4925 | -86.1361
5.4928 | -86.1366
5.4931 | -86.137
5.4934 | -86.1379
5.4937 | -86.1386
5.494 | -86.1387
5.4943 | -86.1389
5.4947 | -86.1392
5.495 | -86.1392
5.4953 | -86.1392
5.4956 | -86.1395
5.4959 | -86.14
5.4962 | -86.1403
5.4965 | -86.1401
5.4968 | -86.1407
5.4971 | -86.1423
5.4974 | -86.1432
5.4978 | -86.143
5.4981 | -86.1434
5.4984 | -86.1443
5.4987 | -86.1451
5.499 | -86.146
5.4993 | -86.1475
5.4996 | -86.1475
5.4999 | -86.1474
5.5002 | -86.1475
5.5006 | -86.1484
5.5009 | -86.1496
5.5012 | -86.1491
5.5015 | -86.1494
5.5018 | -86.15
5.5021 | -86.1506
5.5024 | -86.1519
5.5027 | -86.1531
5.503 | -86.1541
5.5034 | -86.1555
5.5037 | -86.1564
5.504 | -86.1576
5.5043 | -86.1581
5.5046 | -86.1588
5.5049 | -86.1594
5.5052 | -86.1598
5.5055 | -86.1605
5.5058 | -86.1611
5.5061 | -86.1619
5.5065 | -86.1646
5.5068 | -86.1672
5.5071 | -86.1689
5.5074 | -86.17
5.5077 | -86.1711
5.508 | -86.1722
5.5083 | -86.1721
5.5086 | -86.1723
5.5089 | -86.1731
5.5093 | -86.1736
5.5096 | -86.1745
5.5099 | -86.1752
5.5102 | -86.1754
5.5105 | -86.1762
5.5108 | -86.1767
5.5111 | -86.1773
5.5114 | -86.1781
5.5117 | -86.1787
5.512 | -86.1793
5.5124 | -86.1799
5.5127 | -86.1805
5.513 | -86.1814
5.5133 | -86.1828
5.5136 | -86.1836
5.5139 | -86.1844
5.5142 | -86.1851
5.5145 | -86.1857
5.5148 | -86.1863
5.5152 | -86.1881
5.5155 | -86.1899
5.5158 | -86.1909
5.5161 | -86.1928
5.5164 | -86.1945
5.5167 | -86.1953
5.517 | -86.1954
5.5173 | -86.1955
5.5176 | -86.1964
5.5179 | -86.197
5.5183 | -86.1977
5.5186 | -86.1988
5.5189 | -86.2001
5.5192 | -86.2021
5.5195 | -86.2035
5.5198 | -86.2054
5.5201 | -86.2068
5.5204 | -86.2077
5.5207 | -86.2086
5.5211 | -86.2088
5.5214 | -86.2088
5.5217 | -86.2088
5.522 | -86.2094
5.5223 | -86.2102
5.5226 | -86.2114
5.5229 | -86.2127
5.5232 | -86.2141
5.5235 | -86.2144
5.5238 | -86.2147
5.5242 | -86.2157
5.5245 | -86.2171
5.5248 | -86.2184
5.5251 | -86.2191
5.5254 | -86.2201
5.5257 | -86.2207
5.526 | -86.2212
5.5263 | -86.2214
5.5266 | -86.2214
5.527 | -86.2226
5.5273 | -86.2231
5.5276 | -86.2238
5.5279 | -86.2254
5.5282 | -86.2267
5.5285 | -86.2279
5.5288 | -86.2296
5.5291 | -86.2305
5.5294 | -86.2317
5.5297 | -86.2323
5.5301 | -86.2331
5.5304 | -86.2341
5.5307 | -86.235
5.531 | -86.2358
5.5313 | -86.2365
5.5316 | -86.2369
5.5319 | -86.2377
5.5322 | -86.2395
5.5325 | -86.2406
5.5329 | -86.2416
5.5332 | -86.2427
5.5335 | -86.2429
5.5338 | -86.2431
5.5341 | -86.2435
5.5344 | -86.2441
5.5347 | -86.2449
5.535 | -86.2466
5.5353 | -86.2477
5.5356 | -86.2479
5.536 | -86.2467
5.5363 | -86.246
5.5366 | -86.2472
5.5369 | -86.2487
5.5372 | -86.2496
5.5375 | -86.2498
5.5378 | -86.2497
5.5381 | -86.2513
5.5384 | -86.253
5.5388 | -86.2547
5.5391 | -86.2556
5.5394 | -86.2562
5.5397 | -86.2568
5.54 | -86.2582
5.5403 | -86.2601
5.5406 | -86.2608
5.5409 | -86.2621
5.5412 | -86.2632
5.5415 | -86.2637
5.5419 | -86.2645
5.5422 | -86.265
5.5425 | -86.2656
5.5428 | -86.2663
5.5431 | -86.2674
5.5434 | -86.2685
5.5437 | -86.2689
5.544 | -86.2686
5.5443 | -86.269
5.5447 | -86.2699
5.545 | -86.2713
5.5453 | -86.2727
5.5456 | -86.2743
5.5459 | -86.2747
5.5462 | -86.2763
5.5465 | -86.2769
5.5468 | -86.2785
5.5471 | -86.2798
5.5474 | -86.2807
5.5478 | -86.282
5.5481 | -86.283
5.5484 | -86.2836
5.5487 | -86.2836
5.549 | -86.2833
5.5493 | -86.2845
5.5496 | -86.2861
5.5499 | -86.2874
5.5502 | -86.2883
5.5506 | -86.2894
5.5509 | -86.2904
5.5512 | -86.2919
5.5515 | -86.2924
5.5518 | -86.2916
5.5521 | -86.2922
5.5524 | -86.2931
5.5527 | -86.2944
5.553 | -86.2953
5.5533 | -86.296
5.5537 | -86.2971
5.554 | -86.2983
5.5543 | -86.2999
5.5546 | -86.3016
5.5549 | -86.3027
5.5552 | -86.3032
5.5555 | -86.3035
5.5558 | -86.3057
5.5561 | -86.307
5.5565 | -86.308
5.5568 | -86.3106
5.5571 | -86.3132
5.5574 | -86.3146
5.5577 | -86.3155
5.558 | -86.3169
5.5583 | -86.3179
5.5586 | -86.3189
5.5589 | -86.3201
5.5592 | -86.3212
5.5596 | -86.3221
5.5599 | -86.3237
5.5602 | -86.3252
5.5605 | -86.326
5.5608 | -86.3265
5.5611 | -86.3262
5.5614 | -86.3257
5.5617 | -86.3257
5.562 | -86.3266
5.5624 | -86.3284
5.5627 | -86.3304
5.563 | -86.3318
5.5633 | -86.3329
5.5636 | -86.3345
5.5639 | -86.3361
5.5642 | -86.3381
5.5645 | -86.3405
5.5648 | -86.3426
5.5651 | -86.3437
5.5655 | -86.3445
5.5658 | -86.346
5.5661 | -86.3484
5.5664 | -86.3507
5.5667 | -86.3527
5.567 | -86.3537
5.5673 | -86.355
5.5676 | -86.3562
5.5679 | -86.3575
5.5683 | -86.3589
5.5686 | -86.3605
5.5689 | -86.3618
5.5692 | -86.3629
5.5695 | -86.3642
5.5698 | -86.3651
5.5701 | -86.3662
5.5704 | -86.367
5.5707 | -86.3679
5.571 | -86.3686
5.5714 | -86.3697
5.5717 | -86.3711
5.572 | -86.3728
5.5723 | -86.3739
5.5726 | -86.3752
5.5729 | -86.3763
5.5732 | -86.3778
5.5735 | -86.3793
5.5738 | -86.3811
5.5742 | -86.3822
5.5745 | -86.3831
5.5748 | -86.3846
5.5751 | -86.387
5.5754 | -86.3882
5.5757 | -86.3892
5.576 | -86.3899
5.5763 | -86.3909
5.5766 | -86.3927
5.5769 | -86.3942
5.5773 | -86.3956
5.5776 | -86.3981
5.5779 | -86.4011
5.5782 | -86.4033
5.5785 | -86.4049
5.5788 | -86.4069
5.5791 | -86.4088
5.5794 | -86.4112
5.5797 | -86.415
5.5801 | -86.4173
5.5804 | -86.4184
5.5807 | -86.4198
5.581 | -86.4208
5.5813 | -86.4222
5.5816 | -86.4237
5.5819 | -86.4257
5.5822 | -86.4273
5.5825 | -86.428
5.5828 | -86.4292
5.5832 | -86.4301
5.5835 | -86.4312
5.5838 | -86.4325
5.5841 | -86.434
5.5844 | -86.4356
5.5847 | -86.4368
5.585 | -86.4384
5.5853 | -86.4396
5.5856 | -86.4408
5.586 | -86.4423
5.5863 | -86.4435
5.5866 | -86.4447
5.5869 | -86.446
5.5872 | -86.4472
5.5875 | -86.4481
5.5878 | -86.4484
5.5881 | -86.4492
5.5884 | -86.4504
5.5887 | -86.4516
5.5891 | -86.4522
5.5894 | -86.4536
5.5897 | -86.4557
5.59 | -86.4569
5.5903 | -86.4583
5.5906 | -86.4562
5.5909 | -86.4568
5.5912 | -86.4581
5.5915 | -86.4592
5.5919 | -86.4604
5.5922 | -86.4617
5.5925 | -86.4631
5.5928 | -86.4633
5.5931 | -86.4642
5.5934 | -86.4652
5.5937 | -86.4663
5.594 | -86.4676
5.5943 | -86.4695
5.5946 | -86.4704
5.595 | -86.4716
5.5953 | -86.4731
5.5956 | -86.4747
5.5959 | -86.4761
5.5962 | -86.4778
5.5965 | -86.4802
5.5968 | -86.4819
5.5971 | -86.483
5.5974 | -86.4839
5.5978 | -86.4851
5.5981 | -86.4862
5.5984 | -86.4873
5.5987 | -86.4881
5.599 | -86.4893
5.5993 | -86.4906
5.5996 | -86.4926
5.5999 | -86.4936
5.6002 | -86.4948
5.6005 | -86.4965
5.6009 | -86.4979
5.6012 | -86.4992
5.6015 | -86.5008
5.6018 | -86.5019
5.6021 | -86.5022
5.6024 | -86.503
5.6027 | -86.5043
5.603 | -86.5051
5.6033 | -86.5059
5.6037 | -86.507
5.604 | -86.5086
5.6043 | -86.5095
5.6046 | -86.5106
5.6049 | -86.5116
5.6052 | -86.5132
5.6055 | -86.5151
5.6058 | -86.5169
5.6061 | -86.5186
5.6064 | -86.5202
5.6068 | -86.5233
5.6071 | -86.525
5.6074 | -86.5263
5.6077 | -86.5275
5.608 | -86.5284
5.6083 | -86.5304
5.6086 | -86.5323
5.6089 | -86.534
5.6092 | -86.5353
5.6096 | -86.5361
5.6099 | -86.5373
5.6102 | -86.5386
5.6105 | -86.5398
5.6108 | -86.542
5.6111 | -86.5439
5.6114 | -86.5453
5.6117 | -86.5463
5.612 | -86.5471
5.6123 | -86.5484
5.6127 | -86.5502
5.613 | -86.5519
5.6133 | -86.5531
5.6136 | -86.5547
5.6139 | -86.5562
5.6142 | -86.558
5.6145 | -86.5599
5.6148 | -86.5613
5.6151 | -86.5619
5.6155 | -86.5626
5.6158 | -86.5634
5.6161 | -86.5641
5.6164 | -86.5649
5.6167 | -86.5657
5.617 | -86.5662
5.6173 | -86.5676
5.6176 | -86.5692
5.6179 | -86.5705
5.6182 | -86.5719
5.6186 | -86.5727
5.6189 | -86.5736
5.6192 | -86.5747
5.6195 | -86.5756
5.6198 | -86.5766
5.6201 | -86.5768
5.6204 | -86.5776
5.6207 | -86.5784
5.621 | -86.5785
5.6214 | -86.5792
5.6217 | -86.581
5.622 | -86.583
5.6223 | -86.5844
5.6226 | -86.5857
5.6229 | -86.5867
5.6232 | -86.5878
5.6235 | -86.5881
5.6238 | -86.5892
5.6242 | -86.5909
5.6245 | -86.5924
5.6248 | -86.5942
5.6251 | -86.5963
5.6254 | -86.598
5.6257 | -86.5992
5.626 | -86.6006
5.6263 | -86.6028
5.6266 | -86.6029
5.6269 | -86.604
5.6273 | -86.6042
5.6276 | -86.6054
5.6279 | -86.6084
5.6282 | -86.6097
5.6285 | -86.6114
5.6288 | -86.6134
5.6291 | -86.6153
5.6294 | -86.6165
5.6297 | -86.6182
5.63 | -86.6194
5.6304 | -86.6205
5.6307 | -86.6217
5.631 | -86.6229
5.6313 | -86.6238
5.6316 | -86.6252
5.6319 | -86.6264
5.6322 | -86.6281
5.6325 | -86.6297
5.6328 | -86.6314
5.6332 | -86.6335
5.6335 | -86.6346
5.6338 | -86.6361
5.6341 | -86.6373
5.6344 | -86.6384
5.6347 | -86.6394
5.635 | -86.6404
5.6353 | -86.6411
5.6356 | -86.6422
5.636 | -86.6437
5.6363 | -86.6462
5.6366 | -86.6484
5.6369 | -86.6495
5.6372 | -86.6504
5.6375 | -86.6513
5.6378 | -86.6516
5.6381 | -86.6519
5.6384 | -86.6516
5.6387 | -86.6515
5.6391 | -86.6527
5.6394 | -86.6542
5.6397 | -86.655
5.64 | -86.6558
5.6403 | -86.6571
5.6406 | -86.6582
5.6409 | -86.6585
5.6412 | -86.6584
5.6415 | -86.6589
5.6419 | -86.6603
5.6422 | -86.6614
5.6425 | -86.6622
5.6428 | -86.6638
5.6431 | -86.6647
5.6434 | -86.6653
5.6437 | -86.666
5.644 | -86.6665
5.6443 | -86.6675
5.6446 | -86.6675
5.645 | -86.6684
5.6453 | -86.6696
5.6456 | -86.6689
5.6459 | -86.6698
5.6462 | -86.6711
5.6465 | -86.6723
5.6468 | -86.6742
5.6471 | -86.676
5.6474 | -86.676
5.6478 | -86.6769
5.6481 | -86.6782
5.6484 | -86.68
5.6487 | -86.6824
5.649 | -86.684
5.6493 | -86.6846
5.6496 | -86.6856
5.6499 | -86.6872
5.6502 | -86.6868
5.6505 | -86.6862
5.6509 | -86.687
5.6512 | -86.6877
5.6515 | -86.689
5.6518 | -86.6896
5.6521 | -86.6902
5.6524 | -86.6908
5.6527 | -86.6919
5.653 | -86.6932
5.6533 | -86.6922
5.6537 | -86.693
5.654 | -86.6936
5.6543 | -86.6941
5.6546 | -86.695
5.6549 | -86.6965
5.6552 | -86.6972
5.6555 | -86.6969
5.6558 | -86.6976
5.6561 | -86.6985
5.6564 | -86.6997
5.6568 | -86.7001
5.6571 | -86.701
5.6574 | -86.7018
5.6577 | -86.7019
5.658 | -86.703
5.6583 | -86.7047
5.6586 | -86.7056
5.6589 | -86.7054
5.6592 | -86.705
5.6596 | -86.7061
5.6599 | -86.7074
5.6602 | -86.7089
5.6605 | -86.7105
5.6608 | -86.7106
5.6611 | -86.711
5.6614 | -86.7124
5.6617 | -86.7139
5.662 | -86.7146
5.6623 | -86.7153
5.6627 | -86.7164
5.663 | -86.7176
5.6633 | -86.7183
5.6636 | -86.719
5.6639 | -86.7199
5.6642 | -86.7206
5.6645 | -86.7214
5.6648 | -86.7221
5.6651 | -86.7226
5.6655 | -86.7237
5.6658 | -86.7251
5.6661 | -86.7269
5.6664 | -86.7281
5.6667 | -86.7286
5.667 | -86.7284
5.6673 | -86.7305
5.6676 | -86.7323
5.6679 | -86.7329
5.6682 | -86.7346
5.6686 | -86.736
5.6689 | -86.7368
5.6692 | -86.7374
5.6695 | -86.7381
5.6698 | -86.7387
5.6701 | -86.7386
5.6704 | -86.7381
5.6707 | -86.7387
5.671 | -86.7403
5.6714 | -86.7405
5.6717 | -86.7411
5.672 | -86.7417
5.6723 | -86.7425
5.6726 | -86.7447
5.6729 | -86.7458
5.6732 | -86.745
5.6735 | -86.7447
5.6738 | -86.7443
5.6741 | -86.744
5.6745 | -86.7436
5.6748 | -86.7433
5.6751 | -86.7439
5.6754 | -86.7443
5.6757 | -86.7439
5.676 | -86.7447
5.6763 | -86.7461
5.6766 | -86.7471
5.6769 | -86.7477
5.6773 | -86.7487
5.6776 | -86.7504
5.6779 | -86.7527
5.6782 | -86.755
5.6785 | -86.7558
5.6788 | -86.7563
5.6791 | -86.7577
5.6794 | -86.7583
5.6797 | -86.7582
5.68 | -86.7597
5.6804 | -86.7603
5.6807 | -86.7623
5.681 | -86.7632
5.6813 | -86.7633
5.6816 | -86.7626
5.6819 | -86.7629
5.6822 | -86.7632
5.6825 | -86.7635
5.6828 | -86.7644
5.6832 | -86.7637
5.6835 | -86.7625
5.6838 | -86.7626
5.6841 | -86.7629
5.6844 | -86.7633
5.6847 | -86.7636
5.685 | -86.7644
5.6853 | -86.7642
5.6856 | -86.7645
5.6859 | -86.7649
5.6863 | -86.7657
5.6866 | -86.7672
5.6869 | -86.7676
5.6872 | -86.767
5.6875 | -86.7671
5.6878 | -86.7678
5.6881 | -86.7679
5.6884 | -86.7682
5.6887 | -86.7687
5.6891 | -86.7692
5.6894 | -86.7697
5.6897 | -86.7696
5.69 | -86.7699
5.6903 | -86.77
5.6906 | -86.7697
5.6909 | -86.7695
5.6912 | -86.7699
5.6915 | -86.7699
5.6918 | -86.77
5.6922 | -86.7701
5.6925 | -86.77
5.6928 | -86.7709
5.6931 | -86.7711
5.6934 | -86.7712
5.6937 | -86.7721
5.694 | -86.7727
5.6943 | -86.7723
5.6946 | -86.7729
5.695 | -86.7733
5.6953 | -86.7732
5.6956 | -86.7733
5.6959 | -86.7733
5.6962 | -86.7738
5.6965 | -86.7741
5.6968 | -86.7743
5.6971 | -86.7741
5.6974 | -86.7742
5.6977 | -86.7745
5.6981 | -86.7742
5.6984 | -86.7699
5.6987 | -86.7708
5.699 | -86.7708
5.6993 | -86.7707
5.6996 | -86.7707
5.6999 | -86.7721
5.7002 | -86.7731
5.7005 | -86.7732
5.7009 | -86.7732
5.7012 | -86.7731
5.7015 | -86.7731
5.7018 | -86.7732
5.7021 | -86.7737
5.7024 | -86.7737
5.7027 | -86.774
5.703 | -86.774
5.7033 | -86.7735
5.7036 | -86.7732
5.704 | -86.7728
5.7043 | -86.7725
5.7046 | -86.7722
5.7049 | -86.772
5.7052 | -86.7717
5.7055 | -86.7716
5.7058 | -86.7727
5.7061 | -86.7735
5.7064 | -86.774
5.7068 | -86.7741
5.7071 | -86.7744
5.7074 | -86.7733
5.7077 | -86.7727
5.708 | -86.7733
5.7083 | -86.7753
5.7086 | -86.7764
5.7089 | -86.7764
5.7092 | -86.7765
5.7095 | -86.7769
5.7099 | -86.777
5.7102 | -86.777
5.7105 | -86.7778
5.7108 | -86.7781
5.7111 | -86.7776
5.7114 | -86.7774
5.7117 | -86.7782
5.712 | -86.7789
5.7123 | -86.7792
5.7127 | -86.7788
5.713 | -86.7776
5.7133 | -86.7769
5.7136 | -86.7778
5.7139 | -86.7793
5.7142 | -86.7795
5.7145 | -86.7795
5.7148 | -86.7797
5.7151 | -86.78
5.7154 | -86.78
5.7158 | -86.7802
5.7161 | -86.7797
5.7164 | -86.7796
5.7167 | -86.7797
5.717 | -86.7799
5.7173 | -86.7798
5.7176 | -86.779
5.7179 | -86.778
5.7182 | -86.7775
5.7186 | -86.7775
5.7189 | -86.7775
5.7192 | -86.7782
5.7195 | -86.7782
5.7198 | -86.7775
5.7201 | -86.7765
5.7204 | -86.7762
5.7207 | -86.776
5.721 | -86.7748
5.7213 | -86.774
5.7217 | -86.7734
5.722 | -86.7737
5.7223 | -86.7744
5.7226 | -86.7746
5.7229 | -86.7745
5.7232 | -86.774
5.7235 | -86.7739
5.7238 | -86.7734
5.7241 | -86.7725
5.7245 | -86.7717
5.7248 | -86.7712
5.7251 | -86.7711
5.7254 | -86.7704
5.7257 | -86.7703
5.726 | -86.7707
5.7263 | -86.7713
5.7266 | -86.7709
5.7269 | -86.7704
5.7272 | -86.77
5.7276 | -86.7696
5.7279 | -86.7692
5.7282 | -86.7692
5.7285 | -86.7702
5.7288 | -86.7703
5.7291 | -86.7705
5.7294 | -86.7704
5.7297 | -86.7701
5.73 | -86.7704
5.7304 | -86.7701
5.7307 | -86.7695
5.731 | -86.7689
5.7313 | -86.768
5.7316 | -86.767
5.7319 | -86.7655
5.7322 | -86.7648
5.7325 | -86.7636
5.7328 | -86.7624
5.7331 | -86.7621
5.7335 | -86.7621
5.7338 | -86.762
5.7341 | -86.7609
5.7344 | -86.7608
5.7347 | -86.7609
5.735 | -86.7603
5.7353 | -86.7596
5.7356 | -86.7595
5.7359 | -86.7589
5.7363 | -86.7584
5.7366 | -86.7577
5.7369 | -86.7567
5.7372 | -86.7563
5.7375 | -86.7563
5.7378 | -86.756
5.7381 | -86.7554
5.7384 | -86.7548
5.7387 | -86.754
5.739 | -86.7533
5.7394 | -86.7529
5.7397 | -86.7523
5.74 | -86.7526
5.7403 | -86.7524
5.7406 | -86.7515
5.7409 | -86.7513
5.7412 | -86.7506
5.7415 | -86.7488
5.7418 | -86.7486
5.7422 | -86.7485
5.7425 | -86.7479
5.7428 | -86.7471
5.7431 | -86.7465
5.7434 | -86.7454
5.7437 | -86.7446
5.744 | -86.7439
5.7443 | -86.7433
5.7446 | -86.7428
5.7449 | -86.7425
5.7453 | -86.7422
5.7456 | -86.7418
5.7459 | -86.7424
5.7462 | -86.742
5.7465 | -86.7414
5.7468 | -86.7408
5.7471 | -86.7403
5.7474 | -86.7405
5.7477 | -86.7404
5.7481 | -86.7399
5.7484 | -86.7392
5.7487 | -86.7386
5.749 | -86.7384
5.7493 | -86.7379
5.7496 | -86.7374
5.7499 | -86.737
5.7502 | -86.7362
5.7505 | -86.7357
5.7508 | -86.7342
5.7512 | -86.7327
5.7515 | -86.7309
5.7518 | -86.7288
5.7521 | -86.7283
5.7524 | -86.7276
5.7527 | -86.7268
5.753 | -86.7259
5.7533 | -86.7248
5.7536 | -86.7237
5.754 | -86.7231
5.7543 | -86.7222
5.7546 | -86.7215
5.7549 | -86.7212
5.7552 | -86.7207
5.7555 | -86.72
5.7558 | -86.7191
5.7561 | -86.7187
5.7564 | -86.7177
5.7568 | -86.7171
5.7571 | -86.7161
5.7574 | -86.7149
5.7577 | -86.7143
5.758 | -86.7141
5.7583 | -86.7133
5.7586 | -86.7128
5.7589 | -86.7122
5.7592 | -86.712
5.7595 | -86.7112
5.7599 | -86.7109
5.7602 | -86.7104
5.7605 | -86.7098
5.7608 | -86.7092
5.7611 | -86.7088
5.7614 | -86.7082
5.7617 | -86.7077
5.762 | -86.7074
5.7623 | -86.7064
5.7626 | -86.7046
5.763 | -86.7039
5.7633 | -86.7029
5.7636 | -86.7026
5.7639 | -86.7026
5.7642 | -86.702
5.7645 | -86.7015
5.7648 | -86.7016
5.7651 | -86.7011
5.7654 | -86.7004
5.7658 | -86.7005
5.7661 | -86.7
5.7664 | -86.6994
5.7667 | -86.6988
5.767 | -86.6984
5.7673 | -86.6991
5.7676 | -86.7
5.7679 | -86.7011
5.7682 | -86.7016
5.7686 | -86.7015
5.7689 | -86.7011
5.7692 | -86.7005
5.7695 | -86.6994
5.7698 | -86.6978
5.7701 | -86.6973
5.7704 | -86.697
5.7707 | -86.6965
5.771 | -86.6958
5.7713 | -86.6948
5.7717 | -86.6937
5.772 | -86.6935
5.7723 | -86.6937
5.7726 | -86.6933
5.7729 | -86.6934
5.7732 | -86.6938
5.7735 | -86.693
5.7738 | -86.6922
5.7741 | -86.6927
5.7744 | -86.6933
5.7748 | -86.6926
5.7751 | -86.6913
5.7754 | -86.6907
5.7757 | -86.6889
5.776 | -86.687
5.7763 | -86.686
5.7766 | -86.6854
5.7769 | -86.6849
5.7772 | -86.6838
5.7776 | -86.6838
5.7779 | -86.6832
5.7782 | -86.6819
5.7785 | -86.6811
5.7788 | -86.6797
5.7791 | -86.6789
5.7794 | -86.6779
5.7797 | -86.677
5.78 | -86.6763
5.7804 | -86.6729
5.7807 | -86.6704
5.781 | -86.6699
5.7813 | -86.6696
5.7816 | -86.6685
5.7819 | -86.6678
5.7822 | -86.6673
5.7825 | -86.6671
5.7828 | -86.6661
5.7831 | -86.6655
5.7835 | -86.6647
5.7838 | -86.664
5.7841 | -86.664
5.7844 | -86.6636
5.7847 | -86.6629
5.785 | -86.6624
5.7853 | -86.6609
5.7856 | -86.6611
5.7859 | -86.661
5.7863 | -86.6604
5.7866 | -86.6592
5.7869 | -86.6587
5.7872 | -86.6571
5.7875 | -86.6562
5.7878 | -86.6556
5.7881 | -86.6552
5.7884 | -86.6542
5.7887 | -86.6535
5.789 | -86.653
5.7894 | -86.653
5.7897 | -86.6524
5.79 | -86.6513
5.7903 | -86.6507
5.7906 | -86.6502
5.7909 | -86.6502
5.7912 | -86.6495
5.7915 | -86.6486
5.7918 | -86.6475
5.7922 | -86.6461
5.7925 | -86.6453
5.7928 | -86.6443
5.7931 | -86.6434
5.7934 | -86.6421
5.7937 | -86.6413
5.794 | -86.6402
5.7943 | -86.6391
5.7946 | -86.6382
5.7949 | -86.6374
5.7953 | -86.6367
5.7956 | -86.6359
5.7959 | -86.6352
5.7962 | -86.6345
5.7965 | -86.6336
5.7968 | -86.6313
5.7971 | -86.631
5.7974 | -86.631
5.7977 | -86.6308
5.7981 | -86.6297
5.7984 | -86.6282
5.7987 | -86.6275
5.799 | -86.626
5.7993 | -86.6252
5.7996 | -86.624
5.7999 | -86.623
5.8002 | -86.6224
5.8005 | -86.6222
5.8008 | -86.6217
5.8012 | -86.6207
5.8015 | -86.6199
5.8018 | -86.6199
5.8021 | -86.6193
5.8024 | -86.6189
5.8027 | -86.6184
5.803 | -86.6177
5.8033 | -86.617
5.8036 | -86.6157
5.804 | -86.6147
5.8043 | -86.6145
5.8046 | -86.6138
5.8049 | -86.6162
5.8052 | -86.6167
5.8055 | -86.6157
5.8058 | -86.6146
5.8061 | -86.6145
5.8064 | -86.6141
5.8067 | -86.6136
5.8071 | -86.6128
5.8074 | -86.6122
5.8077 | -86.6111
5.808 | -86.6103
5.8083 | -86.609
5.8086 | -86.6077
5.8089 | -86.6073
5.8092 | -86.6064
5.8095 | -86.6056
5.8099 | -86.6045
5.8102 | -86.6038
5.8105 | -86.6033
5.8108 | -86.603
5.8111 | -86.6022
5.8114 | -86.6014
5.8117 | -86.6008
5.812 | -86.6002
5.8123 | -86.5995
5.8126 | -86.5985
5.813 | -86.5967
5.8133 | -86.5947
5.8136 | -86.5936
5.8139 | -86.593
5.8142 | -86.5918
5.8145 | -86.5908
5.8148 | -86.5892
5.8151 | -86.5878
5.8154 | -86.5864
5.8158 | -86.5856
5.8161 | -86.5849
5.8164 | -86.5838
5.8167 | -86.5832
5.817 | -86.5825
5.8173 | -86.5814
5.8176 | -86.5804
5.8179 | -86.5795
5.8182 | -86.5786
5.8185 | -86.578
5.8189 | -86.5779
5.8192 | -86.5774
5.8195 | -86.5774
5.8198 | -86.5772
5.8201 | -86.5771
5.8204 | -86.5767
5.8207 | -86.575
5.821 | -86.5739
5.8213 | -86.5729
5.8217 | -86.5718
5.822 | -86.5697
5.8223 | -86.5683
5.8226 | -86.5681
5.8229 | -86.5669
5.8232 | -86.5672
5.8235 | -86.5679
5.8238 | -86.5672
5.8241 | -86.5668
5.8244 | -86.5669
5.8248 | -86.568
5.8251 | -86.5675
5.8254 | -86.566
5.8257 | -86.5627
5.826 | -86.5615
5.8263 | -86.5618
5.8266 | -86.56
5.8269 | -86.5588
5.8272 | -86.5577
5.8276 | -86.5572
5.8279 | -86.5563
5.8282 | -86.5543
5.8285 | -86.5538
5.8288 | -86.5527
5.8291 | -86.553
5.8294 | -86.5531
5.8297 | -86.5529
5.83 | -86.552
5.8303 | -86.5511
5.8307 | -86.5502
5.831 | -86.5496
5.8313 | -86.5483
5.8316 | -86.5463
5.8319 | -86.546
5.8322 | -86.5453
5.8325 | -86.5439
5.8328 | -86.545
5.8331 | -86.545
5.8335 | -86.5439
5.8338 | -86.5432
5.8341 | -86.5422
5.8344 | -86.5417
5.8347 | -86.5415
5.835 | -86.5417
5.8353 | -86.5414
5.8356 | -86.5401
5.8359 | -86.5395
5.8362 | -86.5386
5.8366 | -86.5375
5.8369 | -86.5367
5.8372 | -86.5363
5.8375 | -86.5361
5.8378 | -86.5358
5.8381 | -86.535
5.8384 | -86.5339
5.8387 | -86.5326
5.839 | -86.5318
5.8394 | -86.5305
5.8397 | -86.5287
5.84 | -86.527
5.8403 | -86.5262
5.8406 | -86.5259
5.8409 | -86.5246
5.8412 | -86.5228
5.8415 | -86.5227
5.8418 | -86.5229
5.8421 | -86.5224
5.8425 | -86.521
5.8428 | -86.5199
5.8431 | -86.5185
5.8434 | -86.5181
5.8437 | -86.5181
5.844 | -86.518
5.8443 | -86.5172
5.8446 | -86.5165
5.8449 | -86.5153
5.8453 | -86.5151
5.8456 | -86.5144
5.8459 | -86.5138
5.8462 | -86.5131
5.8465 | -86.5124
5.8468 | -86.5117
5.8471 | -86.5113
5.8474 | -86.511
5.8477 | -86.5105
5.848 | -86.5099
5.8484 | -86.5087
5.8487 | -86.508
5.849 | -86.5077
5.8493 | -86.5067
5.8496 | -86.5052
5.8499 | -86.5038
5.8502 | -86.5041
5.8505 | -86.5037
5.8508 | -86.5029
5.8512 | -86.5012
5.8515 | -86.5005
5.8518 | -86.4996
5.8521 | -86.4988
5.8524 | -86.4972
5.8527 | -86.4956
5.853 | -86.4946
5.8533 | -86.4943
5.8536 | -86.4938
5.8539 | -86.4933
5.8543 | -86.4922
5.8546 | -86.4917
5.8549 | -86.4909
5.8552 | -86.49
5.8555 | -86.4897
5.8558 | -86.4891
5.8561 | -86.4884
5.8564 | -86.4881
5.8567 | -86.4878
5.8571 | -86.4872
5.8574 | -86.4866
5.8577 | -86.4858
5.858 | -86.4855
5.8583 | -86.4847
5.8586 | -86.4845
5.8589 | -86.4841
5.8592 | -86.4832
5.8595 | -86.4829
5.8598 | -86.4822
5.8602 | -86.4816
5.8605 | -86.4814
5.8608 | -86.4816
5.8611 | -86.4811
5.8614 | -86.4805
5.8617 | -86.48
5.862 | -86.4796
5.8623 | -86.4786
5.8626 | -86.4777
5.863 | -86.4777
5.8633 | -86.4766
5.8636 | -86.4754
5.8639 | -86.4747
5.8642 | -86.473
5.8645 | -86.4717
5.8648 | -86.4706
5.8651 | -86.4701
5.8654 | -86.4688
5.8657 | -86.4674
5.8661 | -86.4668
5.8664 | -86.4664
5.8667 | -86.4657
5.867 | -86.465
5.8673 | -86.4643
5.8676 | -86.4632
5.8679 | -86.4616
5.8682 | -86.4606
5.8685 | -86.4604
5.8689 | -86.4596
5.8692 | -86.4588
5.8695 | -86.458
5.8698 | -86.4574
5.8701 | -86.4569
5.8704 | -86.4563
5.8707 | -86.4557
5.871 | -86.4548
5.8713 | -86.4544
5.8716 | -86.4541
5.872 | -86.4536
5.8723 | -86.4528
5.8726 | -86.4518
5.8729 | -86.4517
5.8732 | -86.4513
5.8735 | -86.4507
5.8738 | -86.4503
5.8741 | -86.4503
5.8744 | -86.45
5.8748 | -86.4496
5.8751 | -86.449
5.8754 | -86.4479
5.8757 | -86.4466
5.876 | -86.4457
5.8763 | -86.4445
5.8766 | -86.4438
5.8769 | -86.4443
5.8772 | -86.4448
5.8775 | -86.4435
5.8779 | -86.4427
5.8782 | -86.4423
5.8785 | -86.4412
5.8788 | -86.4407
5.8791 | -86.4413
5.8794 | -86.4406
5.8797 | -86.4397
5.88 | -86.4392
5.8803 | -86.439
5.8807 | -86.4388
5.881 | -86.4381
5.8813 | -86.4377
5.8816 | -86.4363
5.8819 | -86.4356
5.8822 | -86.4343
5.8825 | -86.433
5.8828 | -86.4306
5.8831 | -86.4276
5.8834 | -86.426
5.8838 | -86.4258
5.8841 | -86.4251
5.8844 | -86.4245
5.8847 | -86.424
5.885 | -86.4236
5.8853 | -86.4232
5.8856 | -86.4226
5.8859 | -86.4223
5.8862 | -86.4216
5.8866 | -86.4205
5.8869 | -86.4191
5.8872 | -86.4182
5.8875 | -86.4174
5.8878 | -86.4169
5.8881 | -86.4159
5.8884 | -86.4154
5.8887 | -86.4144
5.889 | -86.4141
5.8893 | -86.4138
5.8897 | -86.4132
5.89 | -86.412
5.8903 | -86.4119
5.8906 | -86.4111
5.8909 | -86.41
5.8912 | -86.4097
5.8915 | -86.4094
5.8918 | -86.4095
5.8921 | -86.409
5.8925 | -86.4083
5.8928 | -86.4076
5.8931 | -86.4072
5.8934 | -86.4064
5.8937 | -86.4056
5.894 | -86.405
5.8943 | -86.4045
5.8946 | -86.404
5.8949 | -86.4039
5.8952 | -86.4025
5.8956 | -86.4026
5.8959 | -86.402
5.8962 | -86.4016
5.8965 | -86.4004
5.8968 | -86.4
5.8971 | -86.4009
5.8974 | -86.4006
5.8977 | -86.4001
5.898 | -86.3995
5.8984 | -86.3992
5.8987 | -86.3983
5.899 | -86.3979
5.8993 | -86.3968
5.8996 | -86.3945
5.8999 | -86.3936
5.9002 | -86.3933
5.9005 | -86.3926
5.9008 | -86.392
5.9012 | -86.3916
5.9015 | -86.3911
5.9018 | -86.39
5.9021 | -86.3897
5.9024 | -86.3889
5.9027 | -86.388
5.903 | -86.3866
5.9033 | -86.3859
5.9036 | -86.3857
5.9039 | -86.3855
5.9043 | -86.3845
5.9046 | -86.3834
5.9049 | -86.3833
5.9052 | -86.3828
5.9055 | -86.3823
5.9058 | -86.3817
5.9061 | -86.3813
5.9064 | -86.3815
5.9067 | -86.3814
5.907 | -86.3804
5.9074 | -86.3803
5.9077 | -86.3802
5.908 | -86.3796
5.9083 | -86.3788
5.9086 | -86.3782
5.9089 | -86.3775
5.9092 | -86.377
5.9095 | -86.376
5.9098 | -86.3742
5.9102 | -86.3727
5.9105 | -86.3713
5.9108 | -86.3694
5.9111 | -86.368
5.9114 | -86.3677
5.9117 | -86.3671
5.912 | -86.3665
5.9123 | -86.366
5.9126 | -86.3654
5.913 | -86.3647
5.9133 | -86.3637
5.9136 | -86.3631
5.9139 | -86.362
5.9142 | -86.3614
5.9145 | -86.3604
5.9148 | -86.3592
5.9151 | -86.3577
5.9154 | -86.3566
5.9157 | -86.3561
5.9161 | -86.3555
5.9164 | -86.3548
5.9167 | -86.3542
5.917 | -86.3535
5.9173 | -86.3524
5.9176 | -86.3517
5.9179 | -86.3515
5.9182 | -86.3513
5.9185 | -86.351
5.9188 | -86.351
5.9192 | -86.3506
5.9195 | -86.3501
5.9198 | -86.3495
5.9201 | -86.3491
5.9204 | -86.3485
5.9207 | -86.3482
5.921 | -86.3478
5.9213 | -86.3475
5.9216 | -86.3461
5.922 | -86.3453
5.9223 | -86.3451
5.9226 | -86.3449
5.9229 | -86.3446
5.9232 | -86.3443
5.9235 | -86.3442
5.9238 | -86.3442
5.9241 | -86.3447
5.9244 | -86.3447
5.9248 | -86.3446
5.9251 | -86.3441
5.9254 | -86.3435
5.9257 | -86.3424
5.926 | -86.3415
5.9263 | -86.3406
5.9266 | -86.34
5.9269 | -86.3396
5.9272 | -86.339
5.9275 | -86.3376
5.9279 | -86.3361
5.9282 | -86.3353
5.9285 | -86.3349
5.9288 | -86.3342
5.9291 | -86.3336
5.9294 | -86.3335
5.9297 | -86.3333
5.93 | -86.3331
5.9303 | -86.3325
5.9307 | -86.3316
5.931 | -86.3309
5.9313 | -86.3307
5.9316 | -86.331
5.9319 | -86.3321
5.9322 | -86.3322
5.9325 | -86.3321
5.9328 | -86.3321
5.9331 | -86.3322
5.9334 | -86.332
5.9338 | -86.3318
5.9341 | -86.3317
5.9344 | -86.3309
5.9347 | -86.3303
5.935 | -86.3296
5.9353 | -86.3293
5.9356 | -86.329
5.9359 | -86.3284
5.9362 | -86.3284
5.9366 | -86.3277
5.9369 | -86.327
5.9372 | -86.3265
5.9375 | -86.3262
5.9378 | -86.3259
5.9381 | -86.3256
5.9384 | -86.3256
5.9387 | -86.3249
5.939 | -86.3245
5.9393 | -86.3242
5.9397 | -86.3238
5.94 | -86.323
5.9403 | -86.322
5.9406 | -86.3212
5.9409 | -86.3204
5.9412 | -86.3197
5.9415 | -86.3196
5.9418 | -86.3193
5.9421 | -86.3189
5.9425 | -86.319
5.9428 | -86.3197
5.9431 | -86.3192
5.9434 | -86.3172
5.9437 | -86.3157
5.944 | -86.3152
5.9443 | -86.3148
5.9446 | -86.3147
5.9449 | -86.3151
5.9452 | -86.3147
5.9456 | -86.3146
5.9459 | -86.3146
5.9462 | -86.314
5.9465 | -86.3136
5.9468 | -86.3133
5.9471 | -86.313
5.9474 | -86.3128
5.9477 | -86.3131
5.948 | -86.3136
5.9484 | -86.3133
5.9487 | -86.313
5.949 | -86.3124
5.9493 | -86.3124
5.9496 | -86.3113
5.9499 | -86.3101
5.9502 | -86.3083
5.9505 | -86.3075
5.9508 | -86.3074
5.9511 | -86.3075
5.9515 | -86.307
5.9518 | -86.3058
5.9521 | -86.3047
5.9524 | -86.3045
5.9527 | -86.3041
5.953 | -86.3033
5.9533 | -86.3026
5.9536 | -86.302
5.9539 | -86.3013
5.9543 | -86.3002
5.9546 | -86.2996
5.9549 | -86.2994
5.9552 | -86.2997
5.9555 | -86.2982
5.9558 | -86.2964
5.9561 | -86.2958
5.9564 | -86.2951
5.9567 | -86.2946
5.957 | -86.2941
5.9574 | -86.2937
5.9577 | -86.2934
5.958 | -86.2935
5.9583 | -86.2934
5.9586 | -86.293
5.9589 | -86.2925
5.9592 | -86.2919
5.9595 | -86.2914
5.9598 | -86.2908
5.9602 | -86.2902
5.9605 | -86.2894
5.9608 | -86.2894
5.9611 | -86.2897
5.9614 | -86.2906
5.9617 | -86.2912
5.962 | -86.2903
5.9623 | -86.2894
5.9626 | -86.2887
5.9629 | -86.2882
5.9633 | -86.2878
5.9636 | -86.2854
5.9639 | -86.2836
5.9642 | -86.2829
5.9645 | -86.2822
5.9648 | -86.2811
5.9651 | -86.2796
5.9654 | -86.2786
5.9657 | -86.2782
5.9661 | -86.2777
5.9664 | -86.2767
5.9667 | -86.2763
5.967 | -86.2762
5.9673 | -86.2757
5.9676 | -86.2754
5.9679 | -86.2744
5.9682 | -86.2733
5.9685 | -86.2725
5.9688 | -86.2719
5.9692 | -86.271
5.9695 | -86.2696
5.9698 | -86.2682
5.9701 | -86.2676
5.9704 | -86.2675
5.9707 | -86.2666
5.971 | -86.2665
5.9713 | -86.2664
5.9716 | -86.2658
5.972 | -86.2651
5.9723 | -86.2642
5.9726 | -86.2632
5.9729 | -86.2623
5.9732 | -86.2615
5.9735 | -86.2612
5.9738 | -86.2612
5.9741 | -86.2606
5.9744 | -86.26
5.9747 | -86.2596
5.9751 | -86.2595
5.9754 | -86.2593
5.9757 | -86.2586
5.976 | -86.2588
5.9763 | -86.2582
5.9766 | -86.2575
5.9769 | -86.2565
5.9772 | -86.2556
5.9775 | -86.2553
5.9779 | -86.2546
5.9782 | -86.2538
5.9785 | -86.2534
5.9788 | -86.2524
5.9791 | -86.2506
5.9794 | -86.2498
5.9797 | -86.2492
5.98 | -86.2477
5.9803 | -86.2465
5.9806 | -86.2457
5.981 | -86.2451
5.9813 | -86.2445
5.9816 | -86.2428
5.9819 | -86.2421
5.9822 | -86.2417
5.9825 | -86.2411
5.9828 | -86.2407
5.9831 | -86.2403
5.9834 | -86.2399
5.9838 | -86.2394
5.9841 | -86.2394
5.9844 | -86.2388
5.9847 | -86.2383
5.985 | -86.2374
5.9853 | -86.2367
5.9856 | -86.236
5.9859 | -86.2352
5.9862 | -86.2347
5.9865 | -86.2342
5.9869 | -86.2334
5.9872 | -86.2323
5.9875 | -86.2318
5.9878 | -86.2311
5.9881 | -86.2306
5.9884 | -86.2297
5.9887 | -86.2283
5.989 | -86.2272
5.9893 | -86.227
5.9897 | -86.2276
5.99 | -86.2277
5.9903 | -86.2274
5.9906 | -86.2272
5.9909 | -86.2274
5.9912 | -86.227
5.9915 | -86.2265
5.9918 | -86.2263
5.9921 | -86.2259
5.9924 | -86.2269
5.9928 | -86.2286
5.9931 | -86.2287
5.9934 | -86.2283
5.9937 | -86.2277
5.994 | -86.2267
5.9943 | -86.2261
5.9946 | -86.226
5.9949 | -86.2261
5.9952 | -86.2268
5.9956 | -86.2266
5.9959 | -86.2261
5.9962 | -86.226
5.9965 | -86.226
5.9968 | -86.2258
5.9971 | -86.2259
5.9974 | -86.2253
5.9977 | -86.2248
5.998 | -86.2243
5.9983 | -86.2233
5.9987 | -86.2224
5.999 | -86.2221
5.9993 | -86.2212
5.9996 | -86.2207
5.9999 | -86.2204
6.0002 | -86.2198
6.0005 | -86.2195
6.0008 | -86.2195
6.0011 | -86.2193
6.0015 | -86.2185
6.0018 | -86.2181
6.0021 | -86.2174
6.0024 | -86.2167
6.0027 | -86.2164
6.003 | -86.2165
6.0033 | -86.2159
6.0036 | -86.2155
6.0039 | -86.2147
6.0042 | -86.2141
6.0046 | -86.214
6.0049 | -86.214
6.0052 | -86.2144
6.0055 | -86.2145
6.0058 | -86.2144
6.0061 | -86.2146
6.0064 | -86.2143
6.0067 | -86.214
6.007 | -86.2135
6.0074 | -86.2134
6.0077 | -86.2127
6.008 | -86.2114
6.0083 | -86.2097
6.0086 | -86.2078
6.0089 | -86.2064
6.0092 | -86.2051
6.0095 | -86.2038
6.0098 | -86.2035
6.0101 | -86.2031
6.0105 | -86.2026
6.0108 | -86.2021
6.0111 | -86.202
6.0114 | -86.2017
6.0117 | -86.2015
6.012 | -86.2005
6.0123 | -86.2
6.0126 | -86.1996
6.0129 | -86.1994
6.0133 | -86.1995
6.0136 | -86.1995
6.0139 | -86.1989
6.0142 | -86.1984
6.0145 | -86.1972
6.0148 | -86.1967
6.0151 | -86.1964
6.0154 | -86.196
6.0157 | -86.1957
6.016 | -86.1949
6.0164 | -86.194
6.0167 | -86.1929
6.017 | -86.1924
6.0173 | -86.1919
6.0176 | -86.1912
6.0179 | -86.1907
6.0182 | -86.1901
6.0185 | -86.19
6.0188 | -86.1904
6.0192 | -86.1903
6.0195 | -86.1899
6.0198 | -86.1901
6.0201 | -86.1902
6.0204 | -86.1904
6.0207 | -86.19
6.021 | -86.1907
6.0213 | -86.1904
6.0216 | -86.1898
6.0219 | -86.189
6.0223 | -86.1878
6.0226 | -86.1871
6.0229 | -86.1873
6.0232 | -86.1873
6.0235 | -86.187
6.0238 | -86.1864
6.0241 | -86.1862
6.0244 | -86.1862
6.0247 | -86.1865
6.0251 | -86.1867
6.0254 | -86.1871
6.0257 | -86.1876
6.026 | -86.1874
6.0263 | -86.1872
6.0266 | -86.1867
6.0269 | -86.1863
6.0272 | -86.1858
6.0275 | -86.1857
6.0278 | -86.1861
6.0282 | -86.185
6.0285 | -86.1841
6.0288 | -86.1838
6.0291 | -86.1832
6.0294 | -86.1824
6.0297 | -86.1829
6.03 | -86.1827
6.0303 | -86.1823
6.0306 | -86.1815
6.031 | -86.1811
6.0313 | -86.1814
6.0316 | -86.1813
6.0319 | -86.1812
6.0322 | -86.1821
6.0325 | -86.1823
6.0328 | -86.1821
6.0331 | -86.1815
6.0334 | -86.1809
6.0337 | -86.1805
6.0341 | -86.1802
6.0344 | -86.1797
6.0347 | -86.1792
6.035 | -86.1785
6.0353 | -86.1785
6.0356 | -86.1787
6.0359 | -86.1783
6.0362 | -86.1778
6.0365 | -86.1774
6.0369 | -86.1764
6.0372 | -86.1754
6.0375 | -86.1744
6.0378 | -86.1741
6.0381 | -86.1739
6.0384 | -86.1738
6.0387 | -86.1742
6.039 | -86.1733
6.0393 | -86.1734
6.0396 | -86.1735
6.04 | -86.173
6.0403 | -86.1728
6.0406 | -86.1729
6.0409 | -86.1721
6.0412 | -86.1721
6.0415 | -86.1718
6.0418 | -86.1713
6.0421 | -86.1706
6.0424 | -86.1701
6.0428 | -86.1702
6.0431 | -86.1698
6.0434 | -86.1696
6.0437 | -86.1689
6.044 | -86.1694
6.0443 | -86.1703
6.0446 | -86.1701
6.0449 | -86.1699
6.0452 | -86.1701
6.0456 | -86.1703
6.0459 | -86.1698
6.0462 | -86.1688
6.0465 | -86.1691
6.0468 | -86.1702
6.0471 | -86.1706
6.0474 | -86.1704
6.0477 | -86.171
6.048 | -86.1713
6.0483 | -86.1715
6.0487 | -86.1712
6.049 | -86.1713
6.0493 | -86.1712
6.0496 | -86.1711
6.0499 | -86.1715
6.0502 | -86.1713
6.0505 | -86.1708
6.0508 | -86.1713
6.0511 | -86.171
6.0514 | -86.1708
6.0518 | -86.1719
6.0521 | -86.172
6.0524 | -86.1717
6.0527 | -86.1713
6.053 | -86.1711
6.0533 | -86.1712
6.0536 | -86.1722
6.0539 | -86.1721
6.0542 | -86.1718
6.0546 | -86.1718
6.0549 | -86.1719
6.0552 | -86.172
6.0555 | -86.1726
6.0558 | -86.173
6.0561 | -86.1731
6.0564 | -86.1744
6.0567 | -86.175
6.057 | -86.1759
6.0573 | -86.176
6.0577 | -86.176
6.058 | -86.176
6.0583 | -86.1762
6.0586 | -86.1759
6.0589 | -86.1758
6.0592 | -86.1756
6.0595 | -86.1755
6.0598 | -86.1761
6.0601 | -86.177
6.0605 | -86.1768
6.0608 | -86.1764
6.0611 | -86.1766
6.0614 | -86.1763
6.0617 | -86.1763
6.062 | -86.1763
6.0623 | -86.177
6.0626 | -86.1771
6.0629 | -86.1773
6.0633 | -86.1775
6.0636 | -86.178
6.0639 | -86.1781
6.0642 | -86.1785
6.0645 | -86.1786
6.0648 | -86.1786
6.0651 | -86.1787
6.0654 | -86.1785
6.0657 | -86.1793
6.066 | -86.1804
6.0664 | -86.1818
6.0667 | -86.1817
6.067 | -86.1823
6.0673 | -86.1832
6.0676 | -86.1833
6.0679 | -86.184
6.0682 | -86.1845
6.0685 | -86.1848
6.0688 | -86.1848
6.0692 | -86.1849
6.0695 | -86.185
6.0698 | -86.1852
6.0701 | -86.1857
6.0704 | -86.1857
6.0707 | -86.1859
6.071 | -86.1861
6.0713 | -86.1861
6.0716 | -86.186
6.0719 | -86.186
6.0723 | -86.186
6.0726 | -86.1865
6.0729 | -86.187
6.0732 | -86.1873
6.0735 | -86.1872
6.0738 | -86.1869
6.0741 | -86.1873
6.0744 | -86.1877
6.0747 | -86.1888
6.0751 | -86.1894
6.0754 | -86.1894
6.0757 | -86.1894
6.076 | -86.19
6.0763 | -86.1906
6.0766 | -86.191
6.0769 | -86.1911
6.0772 | -86.1913
6.0775 | -86.1914
6.0778 | -86.1923
6.0782 | -86.1931
6.0785 | -86.1928
6.0788 | -86.1932
6.0791 | -86.1933
6.0794 | -86.1936
6.0797 | -86.1936
6.08 | -86.1937
6.0803 | -86.194
6.0806 | -86.1943
6.081 | -86.1954
6.0813 | -86.1977
6.0816 | -86.1986
6.0819 | -86.1993
6.0822 | -86.2002
6.0825 | -86.201
6.0828 | -86.2016
6.0831 | -86.2016
6.0834 | -86.2027
6.0837 | -86.2031
6.0841 | -86.2037
6.0844 | -86.2045
6.0847 | -86.2049
6.085 | -86.2053
6.0853 | -86.2056
6.0856 | -86.2061
6.0859 | -86.2065
6.0862 | -86.2066
6.0865 | -86.2061
6.0869 | -86.2061
6.0872 | -86.2059
6.0875 | -86.2056
6.0878 | -86.2061
6.0881 | -86.2066
6.0884 | -86.2069
6.0887 | -86.2068
6.089 | -86.2063
6.0893 | -86.2053
6.0896 | -86.2043
6.09 | -86.2047
6.0903 | -86.2058
6.0906 | -86.2064
6.0909 | -86.2067
6.0912 | -86.2077
6.0915 | -86.2083
6.0918 | -86.2094
6.0921 | -86.2111
6.0924 | -86.2125
6.0928 | -86.2132
6.0931 | -86.2135
6.0934 | -86.2131
6.0937 | -86.2131
6.094 | -86.2135
6.0943 | -86.2143
6.0946 | -86.2149
6.0949 | -86.2154
6.0952 | -86.2158
6.0955 | -86.2162
6.0959 | -86.2165
6.0962 | -86.2174
6.0965 | -86.2184
6.0968 | -86.2186
6.0971 | -86.219
6.0974 | -86.2194
6.0977 | -86.2183
6.098 | -86.2177
6.0983 | -86.2181
6.0987 | -86.2185
6.099 | -86.2189
6.0993 | -86.2194
6.0996 | -86.2201
6.0999 | -86.2209
6.1002 | -86.2214
6.1005 | -86.2219
6.1008 | -86.2224
6.1011 | -86.2229
6.1014 | -86.2244
6.1018 | -86.2249
6.1021 | -86.2257
6.1024 | -86.2262
6.1027 | -86.2268
6.103 | -86.2275
6.1033 | -86.2282
6.1036 | -86.2287
6.1039 | -86.2294
6.1042 | -86.23
6.1046 | -86.2306
6.1049 | -86.2301
6.1052 | -86.23
6.1055 | -86.2307
6.1058 | -86.2313
6.1061 | -86.2314
6.1064 | -86.2321
6.1067 | -86.232
6.107 | -86.2322
6.1073 | -86.2326
6.1077 | -86.2335
6.108 | -86.2361
6.1083 | -86.2382
6.1086 | -86.2387
6.1089 | -86.239
6.1092 | -86.2393
6.1095 | -86.2399
6.1098 | -86.2403
6.1101 | -86.2406
6.1105 | -86.242
6.1108 | -86.2424
6.1111 | -86.2427
6.1114 | -86.2432
6.1117 | -86.244
6.112 | -86.245
6.1123 | -86.2457
6.1126 | -86.2459
6.1129 | -86.2458
6.1132 | -86.2463
6.1136 | -86.2473
6.1139 | -86.248
6.1142 | -86.2486
6.1145 | -86.2492
6.1148 | -86.2488
6.1151 | -86.2491
6.1154 | -86.25
6.1157 | -86.251
6.116 | -86.2515
6.1164 | -86.2523
6.1167 | -86.2532
6.117 | -86.2542
6.1173 | -86.2551
6.1176 | -86.2559
6.1179 | -86.2577
6.1182 | -86.2583
6.1185 | -86.2586
6.1188 | -86.2592
6.1191 | -86.2596
6.1195 | -86.2609
6.1198 | -86.2614
6.1201 | -86.2623
6.1204 | -86.2629
6.1207 | -86.2622
6.121 | -86.2608
6.1213 | -86.2599
6.1216 | -86.2601
6.1219 | -86.2599
6.1223 | -86.2597
6.1226 | -86.26
6.1229 | -86.2607
6.1232 | -86.2601
6.1235 | -86.2586
6.1238 | -86.2598
6.1241 | -86.2617
6.1244 | -86.2624
6.1247 | -86.263
6.125 | -86.2643
6.1254 | -86.2654
6.1257 | -86.2663
6.126 | -86.2666
6.1263 | -86.2672
6.1266 | -86.2681
6.1269 | -86.2686
6.1272 | -86.2698
6.1275 | -86.2714
6.1278 | -86.2728
6.1282 | -86.2733
6.1285 | -86.274
6.1288 | -86.2725
6.1291 | -86.2727
6.1294 | -86.2742
6.1297 | -86.2753
6.13 | -86.2757
6.1303 | -86.276
6.1306 | -86.2766
6.1309 | -86.2772
6.1313 | -86.2782
6.1316 | -86.2791
6.1319 | -86.28
6.1322 | -86.2811
6.1325 | -86.2824
6.1328 | -86.2837
6.1331 | -86.2848
6.1334 | -86.2855
6.1337 | -86.2863
6.1341 | -86.287
6.1344 | -86.2874
6.1347 | -86.2885
6.135 | -86.2903
6.1353 | -86.2923
6.1356 | -86.2941
6.1359 | -86.295
6.1362 | -86.2956
6.1365 | -86.2963
6.1368 | -86.2969
6.1372 | -86.2984
6.1375 | -86.2998
6.1378 | -86.301
6.1381 | -86.3013
6.1384 | -86.3013
6.1387 | -86.302
6.139 | -86.3028
6.1393 | -86.3029
6.1396 | -86.3027
6.14 | -86.3031
6.1403 | -86.304
6.1406 | -86.3053
6.1409 | -86.306
6.1412 | -86.3061
6.1415 | -86.3066
6.1418 | -86.3077
6.1421 | -86.3082
6.1424 | -86.3086
6.1427 | -86.3091
6.1431 | -86.31
6.1434 | -86.3112
6.1437 | -86.3121
6.144 | -86.3128
6.1443 | -86.3135
6.1446 | -86.3151
6.1449 | -86.3166
6.1452 | -86.317
6.1455 | -86.3162
6.1459 | -86.3167
6.1462 | -86.3181
6.1465 | -86.3197
6.1468 | -86.3206
6.1471 | -86.3215
6.1474 | -86.3215
6.1477 | -86.3223
6.148 | -86.3231
6.1483 | -86.324
6.1486 | -86.3248
6.149 | -86.3252
6.1493 | -86.3269
6.1496 | -86.3277
6.1499 | -86.3283
6.1502 | -86.3279
6.1505 | -86.3275
6.1508 | -86.328
6.1511 | -86.3285
6.1514 | -86.3294
6.1518 | -86.3303
6.1521 | -86.331
6.1524 | -86.3318
6.1527 | -86.3324
6.153 | -86.333
6.1533 | -86.3341
6.1536 | -86.335
6.1539 | -86.3363
6.1542 | -86.3383
6.1545 | -86.3387
6.1549 | -86.3389
6.1552 | -86.3398
6.1555 | -86.3405
6.1558 | -86.3413
6.1561 | -86.3424
6.1564 | -86.3428
6.1567 | -86.3446
6.157 | -86.3459
6.1573 | -86.3466
6.1577 | -86.3474
6.158 | -86.3484
6.1583 | -86.349
6.1586 | -86.3492
6.1589 | -86.349
6.1592 | -86.3486
6.1595 | -86.349
6.1598 | -86.349
6.1601 | -86.3487
6.1604 | -86.3495
6.1608 | -86.3502
6.1611 | -86.3508
6.1614 | -86.3511
6.1617 | -86.3524
6.162 | -86.3534
6.1623 | -86.354
6.1626 | -86.3533
6.1629 | -86.3518
6.1632 | -86.3521
6.1636 | -86.3524
6.1639 | -86.3517
6.1642 | -86.3518
6.1645 | -86.3527
6.1648 | -86.3549
6.1651 | -86.3569
6.1654 | -86.3589
6.1657 | -86.36
6.166 | -86.3606
6.1663 | -86.3606
6.1667 | -86.361
6.167 | -86.3613
6.1673 | -86.3624
6.1676 | -86.3634
6.1679 | -86.3629
6.1682 | -86.3632
6.1685 | -86.3643
6.1688 | -86.3645
6.1691 | -86.3648
6.1695 | -86.3651
6.1698 | -86.3657
6.1701 | -86.3658
6.1704 | -86.3661
6.1707 | -86.3666
6.171 | -86.3681
6.1713 | -86.3691
6.1716 | -86.368
6.1719 | -86.3672
6.1722 | -86.3679
6.1726 | -86.3687
6.1729 | -86.3695
6.1732 | -86.3701
6.1735 | -86.3706
6.1738 | -86.3712
6.1741 | -86.3722
6.1744 | -86.3732
6.1747 | -86.3739
6.175 | -86.3746
6.1754 | -86.3753
6.1757 | -86.3763
6.176 | -86.3765
6.1763 | -86.3764
6.1766 | -86.3773
6.1769 | -86.3782
6.1772 | -86.3785
6.1775 | -86.3785
6.1778 | -86.3792
6.1781 | -86.3789
6.1785 | -86.3792
6.1788 | -86.3796
6.1791 | -86.3808
6.1794 | -86.3816
6.1797 | -86.3813
6.18 | -86.3812
6.1803 | -86.3818
6.1806 | -86.3823
6.1809 | -86.3829
6.1813 | -86.3834
6.1816 | -86.3836
6.1819 | -86.3838
6.1822 | -86.3834
6.1825 | -86.3835
6.1828 | -86.3826
6.1831 | -86.3811
6.1834 | -86.3801
6.1837 | -86.3799
6.1841 | -86.3803
6.1844 | -86.3812
6.1847 | -86.3828
6.185 | -86.3832
6.1853 | -86.3834
6.1856 | -86.3838
6.1859 | -86.3815
6.1862 | -86.3798
6.1865 | -86.3793
6.1868 | -86.3796
6.1872 | -86.3801
6.1875 | -86.3802
6.1878 | -86.3805
6.1881 | -86.3808
6.1884 | -86.3807
6.1887 | -86.3808
6.189 | -86.381
6.1893 | -86.3813
6.1896 | -86.3799
6.1899 | -86.3776
6.1903 | -86.3779
6.1906 | -86.377
6.1909 | -86.3768
6.1912 | -86.3757
6.1915 | -86.3767
6.1918 | -86.3772
6.1921 | -86.3776
6.1924 | -86.3779
6.1927 | -86.3786
6.1931 | -86.3789
6.1934 | -86.3791
6.1937 | -86.3794
6.194 | -86.3797
6.1943 | -86.3802
6.1946 | -86.3812
6.1949 | -86.3807
6.1952 | -86.3807
6.1955 | -86.3813
6.1959 | -86.3808
6.1962 | -86.3824
6.1965 | -86.384
6.1968 | -86.3842
6.1971 | -86.3844
6.1974 | -86.385
6.1977 | -86.3846
6.198 | -86.3845
6.1983 | -86.3848
6.1986 | -86.3851
6.199 | -86.3857
6.1993 | -86.3864
6.1996 | -86.3867
6.1999 | -86.3855
6.2002 | -86.3849
6.2005 | -86.3847
6.2008 | -86.3849
6.2011 | -86.3852
6.2014 | -86.3863
6.2017 | -86.3863
6.2021 | -86.3867
6.2024 | -86.3869
6.2027 | -86.3875
6.203 | -86.3877
6.2033 | -86.3881
6.2036 | -86.388
6.2039 | -86.3884
6.2042 | -86.3881
6.2045 | -86.387
6.2049 | -86.3856
6.2052 | -86.3854
6.2055 | -86.3854
6.2058 | -86.3854
6.2061 | -86.3856
6.2064 | -86.3857
6.2067 | -86.3847
6.207 | -86.3848
6.2073 | -86.3841
6.2077 | -86.3839
6.208 | -86.3839
6.2083 | -86.3842
6.2086 | -86.3843
6.2089 | -86.3783
6.2092 | -86.3785
6.2095 | -86.3786
6.2098 | -86.3786
6.2101 | -86.3791
6.2104 | -86.379
6.2108 | -86.3795
6.2111 | -86.3796
6.2114 | -86.3798
6.2117 | -86.3797
6.212 | -86.3796
6.2123 | -86.3791
6.2126 | -86.3784
6.2129 | -86.3779
6.2132 | -86.3781
6.2136 | -86.3789
6.2139 | -86.3798
6.2142 | -86.3796
6.2145 | -86.3787
6.2148 | -86.3778
6.2151 | -86.3775
6.2154 | -86.3776
6.2157 | -86.3776
6.216 | -86.3781
6.2163 | -86.3782
6.2167 | -86.3786
6.217 | -86.3782
6.2173 | -86.3776
6.2176 | -86.3772
6.2179 | -86.3743
6.2182 | -86.374
6.2185 | -86.3746
6.2188 | -86.3742
6.2191 | -86.3727
6.2195 | -86.3718
6.2198 | -86.3718
6.2201 | -86.3719
6.2204 | -86.3718
6.2207 | -86.3711
6.221 | -86.3712
6.2213 | -86.3716
6.2216 | -86.3711
6.2219 | -86.3709
6.2222 | -86.3706
6.2226 | -86.3708
6.2229 | -86.3706
6.2232 | -86.3693
6.2235 | -86.3689
6.2238 | -86.3683
6.2241 | -86.3679
6.2244 | -86.3661
6.2247 | -86.3654
6.225 | -86.3648
6.2254 | -86.3649
6.2257 | -86.3649
6.226 | -86.3647
6.2263 | -86.3644
6.2266 | -86.3641
6.2269 | -86.3632
6.2272 | -86.3617
6.2275 | -86.3606
6.2278 | -86.3604
6.2281 | -86.3604
6.2285 | -86.3598
6.2288 | -86.3593
6.2291 | -86.3575
6.2294 | -86.3563
6.2297 | -86.3559
6.23 | -86.3554
6.2303 | -86.3546
6.2306 | -86.3541
6.2309 | -86.3544
6.2313 | -86.3548
6.2316 | -86.3537
6.2319 | -86.352
6.2322 | -86.3521
6.2325 | -86.3519
6.2328 | -86.352
6.2331 | -86.3517
6.2334 | -86.351
6.2337 | -86.3492
6.234 | -86.3482
6.2344 | -86.3473
6.2347 | -86.3464
6.235 | -86.3454
6.2353 | -86.3446
6.2356 | -86.3428
6.2359 | -86.3384
6.2362 | -86.3362
6.2365 | -86.3339
6.2368 | -86.3328
6.2372 | -86.3306
6.2375 | -86.3288
6.2378 | -86.3275
6.2381 | -86.3269
6.2384 | -86.326
6.2387 | -86.325
6.239 | -86.3243
6.2393 | -86.3243
6.2396 | -86.3237
6.2399 | -86.323
6.2403 | -86.3224
6.2406 | -86.3225
6.2409 | -86.3219
6.2412 | -86.3202
6.2415 | -86.3196
6.2418 | -86.3186
6.2421 | -86.3177
6.2424 | -86.3171
6.2427 | -86.3158
6.2431 | -86.3142
6.2434 | -86.3127
6.2437 | -86.3115
6.244 | -86.3111
6.2443 | -86.3111
6.2446 | -86.3124
6.2449 | -86.3105
6.2452 | -86.3098
6.2455 | -86.3086
6.2458 | -86.3082
6.2462 | -86.3075
6.2465 | -86.3068
6.2468 | -86.3057
6.2471 | -86.3045
6.2474 | -86.3036
6.2477 | -86.3027
6.248 | -86.3016
6.2483 | -86.2994
6.2486 | -86.2974
6.249 | -86.2956
6.2493 | -86.2944
6.2496 | -86.2935
6.2499 | -86.292
6.2502 | -86.2905
6.2505 | -86.2896
6.2508 | -86.2889
6.2511 | -86.2873
6.2514 | -86.2855
6.2517 | -86.2848
6.2521 | -86.2835
6.2524 | -86.2817
6.2527 | -86.28
6.253 | -86.2782
6.2533 | -86.2764
6.2536 | -86.2751
6.2539 | -86.273
6.2542 | -86.2716
6.2545 | -86.2699
6.2549 | -86.2685
6.2552 | -86.2673
6.2555 | -86.2664
6.2558 | -86.2649
6.2561 | -86.2636
6.2564 | -86.2626
6.2567 | -86.2614
6.257 | -86.2601
6.2573 | -86.2587
6.2576 | -86.2573
6.258 | -86.2554
6.2583 | -86.2523
6.2586 | -86.2498
6.2589 | -86.248
6.2592 | -86.2476
6.2595 | -86.2479
6.2598 | -86.2472
6.2601 | -86.2454
6.2604 | -86.2433
6.2608 | -86.2421
6.2611 | -86.2417
6.2614 | -86.2403
6.2617 | -86.2385
6.262 | -86.2368
6.2623 | -86.2353
6.2626 | -86.2318
6.2629 | -86.2305
6.2632 | -86.2293
6.2635 | -86.2278
6.2639 | -86.2263
6.2642 | -86.2248
6.2645 | -86.223
6.2648 | -86.2212
6.2651 | -86.2195
6.2654 | -86.2178
6.2657 | -86.2162
6.266 | -86.2146
6.2663 | -86.213
6.2667 | -86.2116
6.267 | -86.2096
6.2673 | -86.2078
6.2676 | -86.2063
6.2679 | -86.2047
6.2682 | -86.2028
6.2685 | -86.2008
6.2688 | -86.1988
6.2691 | -86.1969
6.2694 | -86.195
6.2698 | -86.1937
6.2701 | -86.1927
6.2704 | -86.1919
6.2707 | -86.1896
6.271 | -86.1882
6.2713 | -86.1862
6.2716 | -86.1823
6.2719 | -86.1809
6.2722 | -86.18
6.2726 | -86.1781
6.2729 | -86.1767
6.2732 | -86.1749
6.2735 | -86.1737
6.2738 | -86.1725
6.2741 | -86.1705
6.2744 | -86.1681
6.2747 | -86.1645
6.275 | -86.1603
6.2753 | -86.1581
6.2757 | -86.1558
6.276 | -86.1534
6.2763 | -86.1513
6.2766 | -86.1493
6.2769 | -86.1473
6.2772 | -86.1454
6.2775 | -86.1439
6.2778 | -86.1421
6.2781 | -86.1401
6.2785 | -86.1383
6.2788 | -86.1362
6.2791 | -86.1337
6.2794 | -86.1319
6.2797 | -86.1299
6.28 | -86.1271
6.2803 | -86.1244
6.2806 | -86.1215
6.2809 | -86.1192
6.2812 | -86.1165
6.2816 | -86.1142
6.2819 | -86.1122
6.2822 | -86.1099
6.2825 | -86.1068
6.2828 | -86.1032
6.2831 | -86.1003
6.2834 | -86.0982
6.2837 | -86.0963
6.284 | -86.0937
6.2844 | -86.0911
6.2847 | -86.0882
6.285 | -86.0856
6.2853 | -86.083
6.2856 | -86.0801
6.2859 | -86.0787
6.2862 | -86.0774
6.2865 | -86.0752
6.2868 | -86.0723
6.2871 | -86.0707
6.2875 | -86.0687
6.2878 | -86.0667
6.2881 | -86.0644
6.2884 | -86.062
6.2887 | -86.0595
6.289 | -86.0566
6.2893 | -86.0547
6.2896 | -86.0512
6.2899 | -86.0498
6.2903 | -86.0478
6.2906 | -86.0457
6.2909 | -86.0435
6.2912 | -86.0412
6.2915 | -86.0391
6.2918 | -86.0369
6.2921 | -86.0347
6.2924 | -86.0323
6.2927 | -86.0301
6.293 | -86.028
6.2934 | -86.026
6.2937 | -86.0239
6.294 | -86.0217
6.2943 | -86.019
6.2946 | -86.0164
6.2949 | -86.0145
6.2952 | -86.0124
6.2955 | -86.0096
6.2958 | -86.0068
6.2962 | -86.0046
6.2965 | -86.0023
6.2968 | -85.9996
6.2971 | -85.997
6.2974 | -85.9948
6.2977 | -85.9934
6.298 | -85.9921
6.2983 | -85.9905
6.2986 | -85.9879
6.2989 | -85.9859
6.2993 | -85.9838
6.2996 | -85.9816
6.2999 | -85.9792
6.3002 | -85.9774
6.3005 | -85.9758
6.3008 | -85.9734
6.3011 | -85.9705
6.3014 | -85.9672
6.3017 | -85.9641
6.3021 | -85.9618
6.3024 | -85.9596
6.3027 | -85.9573
6.303 | -85.9553
6.3033 | -85.9531
6.3036 | -85.9507
6.3039 | -85.9479
6.3042 | -85.9452
6.3045 | -85.9431
6.3048 | -85.9407
6.3052 | -85.9391
6.3055 | -85.9376
6.3058 | -85.9367
6.3061 | -85.9362
6.3064 | -85.9349
6.3067 | -85.9326
6.307 | -85.9302
6.3073 | -85.9276
6.3076 | -85.9254
6.308 | -85.9229
6.3083 | -85.9205
6.3086 | -85.9176
6.3089 | -85.914
6.3092 | -85.9112
6.3095 | -85.9087
6.3098 | -85.9061
6.3101 | -85.9043
6.3104 | -85.9027
6.3107 | -85.9006
6.3111 | -85.8977
6.3114 | -85.8955
6.3117 | -85.8932
6.312 | -85.8906
6.3123 | -85.8884
6.3126 | -85.885
6.3129 | -85.8819
6.3132 | -85.8796
6.3135 | -85.878
6.3139 | -85.8755
6.3142 | -85.8731
6.3145 | -85.8713
6.3148 | -85.8687
6.3151 | -85.866
6.3154 | -85.8635
6.3157 | -85.8609
6.316 | -85.858
6.3163 | -85.8541
6.3166 | -85.8512
6.317 | -85.8486
6.3173 | -85.8466
6.3176 | -85.8442
6.3179 | -85.8416
6.3182 | -85.839
6.3185 | -85.8365
6.3188 | -85.8342
6.3191 | -85.8326
6.3194 | -85.8307
6.3198 | -85.828
6.3201 | -85.8257
6.3204 | -85.8224
6.3207 | -85.8195
6.321 | -85.8173
6.3213 | -85.8145
6.3216 | -85.8114
6.3219 | -85.8088
6.3222 | -85.8062
6.3225 | -85.8033
6.3229 | -85.801
6.3232 | -85.7987
6.3235 | -85.7969
6.3238 | -85.7958
6.3241 | -85.7931
6.3244 | -85.7906
6.3247 | -85.7881
6.325 | -85.7854
6.3253 | -85.7821
6.3257 | -85.779
6.326 | -85.7765
6.3263 | -85.7743
6.3266 | -85.7717
6.3269 | -85.7696
6.3272 | -85.7676
6.3275 | -85.7653
6.3278 | -85.7623
6.3281 | -85.7598
6.3285 | -85.7575
6.3288 | -85.7551
6.3291 | -85.7535
6.3294 | -85.7521
6.3297 | -85.7496
6.33 | -85.7471
6.3303 | -85.7446
6.3306 | -85.7422
6.3309 | -85.7399
6.3312 | -85.7374
6.3316 | -85.7347
6.3319 | -85.7323
6.3322 | -85.73
6.3325 | -85.7279
6.3328 | -85.7255
6.3331 | -85.7232
6.3334 | -85.7209
6.3337 | -85.7187
6.334 | -85.7163
6.3343 | -85.7141
6.3347 | -85.7117
6.335 | -85.7104
6.3353 | -85.709
6.3356 | -85.7076
6.3359 | -85.7066
6.3362 | -85.705
6.3365 | -85.7026
6.3368 | -85.7004
6.3371 | -85.6981
6.3375 | -85.6956
6.3378 | -85.6931
6.3381 | -85.6911
6.3384 | -85.6893
6.3387 | -85.687
6.339 | -85.6842
6.3393 | -85.6828
6.3396 | -85.6811
6.3399 | -85.6792
6.3403 | -85.6774
6.3406 | -85.6758
6.3409 | -85.6735
6.3412 | -85.6708
6.3415 | -85.6684
6.3418 | -85.6658
6.3421 | -85.6629
6.3424 | -85.6607
6.3427 | -85.6583
6.343 | -85.6552
6.3434 | -85.6535
6.3437 | -85.6508
6.344 | -85.6486
6.3443 | -85.6468
6.3446 | -85.6448
6.3449 | -85.6427
6.3452 | -85.6403
6.3455 | -85.6379
6.3458 | -85.6356
6.3461 | -85.6331
6.3465 | -85.631
6.3468 | -85.6291
6.3471 | -85.6268
6.3474 | -85.6251
6.3477 | -85.6231
6.348 | -85.6204
6.3483 | -85.6178
6.3486 | -85.6154
6.3489 | -85.6128
6.3493 | -85.6106
6.3496 | -85.6085
6.3499 | -85.6064
6.3502 | -85.6035
6.3505 | -85.6011
6.3508 | -85.5989
6.3511 | -85.5966
6.3514 | -85.5947
6.3517 | -85.5929
6.3521 | -85.5907
6.3524 | -85.5895
6.3527 | -85.5876
6.353 | -85.5857
6.3533 | -85.5835
6.3536 | -85.5816
6.3539 | -85.5805
6.3542 | -85.5787
6.3545 | -85.5761
6.3548 | -85.5735
6.3552 | -85.5715
6.3555 | -85.5692
6.3558 | -85.567
6.3561 | -85.5651
6.3564 | -85.5627
6.3567 | -85.5605
6.357 | -85.5586
6.3573 | -85.5566
6.3576 | -85.5545
6.358 | -85.5525
6.3583 | -85.5503
6.3586 | -85.5481
6.3589 | -85.5458
6.3592 | -85.5427
6.3595 | -85.5395
6.3598 | -85.5368
6.3601 | -85.5346
6.3604 | -85.5331
6.3607 | -85.5329
6.3611 | -85.5324
6.3614 | -85.5306
6.3617 | -85.5288
6.362 | -85.5269
6.3623 | -85.525
6.3626 | -85.5229
6.3629 | -85.5214
6.3632 | -85.5199
6.3635 | -85.518
6.3639 | -85.5162
6.3642 | -85.5141
6.3645 | -85.5119
6.3648 | -85.5093
6.3651 | -85.5074
6.3654 | -85.5053
6.3657 | -85.5034
6.366 | -85.5014
6.3663 | -85.4995
6.3666 | -85.4976
6.367 | -85.4962
6.3673 | -85.4945
6.3676 | -85.4927
6.3679 | -85.4911
6.3682 | -85.4896
6.3685 | -85.4881
6.3688 | -85.4863
6.3691 | -85.4844
6.3694 | -85.4825
6.3698 | -85.4806
6.3701 | -85.4789
6.3704 | -85.4785
6.3707 | -85.4772
6.371 | -85.4764
6.3713 | -85.4753
6.3716 | -85.4738
6.3719 | -85.4723
6.3722 | -85.4705
6.3725 | -85.4687
6.3729 | -85.4669
6.3732 | -85.4653
6.3735 | -85.4635
6.3738 | -85.4618
6.3741 | -85.4602
6.3744 | -85.4585
6.3747 | -85.4567
6.375 | -85.455
6.3753 | -85.4534
6.3757 | -85.4519
6.376 | -85.4503
6.3763 | -85.4486
6.3766 | -85.4469
6.3769 | -85.4453
6.3772 | -85.4439
6.3775 | -85.4428
6.3778 | -85.4415
6.3781 | -85.4406
6.3784 | -85.439
6.3788 | -85.4376
6.3791 | -85.4372
6.3794 | -85.4364
6.3797 | -85.4354
6.38 | -85.4345
6.3803 | -85.4333
6.3806 | -85.4316
6.3809 | -85.4299
6.3812 | -85.4286
6.3816 | -85.427
6.3819 | -85.4251
6.3822 | -85.4237
6.3825 | -85.4228
6.3828 | -85.4218
6.3831 | -85.4205
6.3834 | -85.4193
6.3837 | -85.4186
6.384 | -85.4181
6.3843 | -85.4169
6.3847 | -85.4157
6.385 | -85.4144
6.3853 | -85.413
6.3856 | -85.4116
6.3859 | -85.4103
6.3862 | -85.409
6.3865 | -85.4076
6.3868 | -85.4061
6.3871 | -85.4046
6.3875 | -85.4029
6.3878 | -85.4014
6.3881 | -85.4016
6.3884 | -85.4002
6.3887 | -85.3987
6.389 | -85.3975
6.3893 | -85.3967
6.3896 | -85.395
6.3899 | -85.3937
6.3902 | -85.3927
6.3906 | -85.3912
6.3909 | -85.3899
6.3912 | -85.3886
6.3915 | -85.387
6.3918 | -85.3857
6.3921 | -85.3847
6.3924 | -85.3841
6.3927 | -85.3835
6.393 | -85.3824
6.3934 | -85.3812
6.3937 | -85.3805
6.394 | -85.3799
6.3943 | -85.3786
6.3946 | -85.3771
6.3949 | -85.3758
6.3952 | -85.3746
6.3955 | -85.3733
6.3958 | -85.372
6.3961 | -85.3708
6.3965 | -85.37
6.3968 | -85.369
6.3971 | -85.3694
6.3974 | -85.3687
6.3977 | -85.368
6.398 | -85.3669
6.3983 | -85.3659
6.3986 | -85.3653
6.3989 | -85.3651
6.3993 | -85.3645
6.3996 | -85.3633
6.3999 | -85.3624
6.4002 | -85.3616
6.4005 | -85.3603
6.4008 | -85.3596
6.4011 | -85.3592
6.4014 | -85.3585
6.4017 | -85.3576
6.402 | -85.3569
6.4024 | -85.3564
6.4027 | -85.3557
6.403 | -85.3547
6.4033 | -85.3539
6.4036 | -85.3531
6.4039 | -85.3526
6.4042 | -85.3514
6.4045 | -85.3505
6.4048 | -85.3497
6.4052 | -85.3493
6.4055 | -85.3485
6.4058 | -85.3478
6.4061 | -85.3478
6.4064 | -85.347
6.4067 | -85.3463
6.407 | -85.346
6.4073 | -85.3459
6.4076 | -85.3455
6.4079 | -85.3448
6.4083 | -85.3439
6.4086 | -85.343
6.4089 | -85.3422
6.4092 | -85.3423
6.4095 | -85.3411
6.4098 | -85.3398
6.4101 | -85.34
6.4104 | -85.3398
6.4107 | -85.339
6.4111 | -85.3378
6.4114 | -85.3369
6.4117 | -85.3372
6.412 | -85.337
6.4123 | -85.3366
6.4126 | -85.3362
6.4129 | -85.3354
6.4132 | -85.3346
6.4135 | -85.334
6.4138 | -85.3336
6.4142 | -85.3332
6.4145 | -85.333
6.4148 | -85.3327
6.4151 | -85.3339
6.4154 | -85.3338
6.4157 | -85.3333
6.416 | -85.3332
6.4163 | -85.3329
6.4166 | -85.3327
6.417 | -85.3324
6.4173 | -85.3321
6.4176 | -85.3322
6.4179 | -85.3323
6.4182 | -85.3322
6.4185 | -85.3321
6.4188 | -85.3319
6.4191 | -85.3318
6.4194 | -85.3321
6.4197 | -85.3323
6.4201 | -85.3316
6.4204 | -85.3311
6.4207 | -85.3308
6.421 | -85.3304
6.4213 | -85.33
6.4216 | -85.3298
6.4219 | -85.3298
6.4222 | -85.3296
6.4225 | -85.3292
6.4229 | -85.329
6.4232 | -85.3288
6.4235 | -85.3287
6.4238 | -85.3287
6.4241 | -85.3305
6.4244 | -85.3304
6.4247 | -85.3304
6.425 | -85.3305
6.4253 | -85.3311
6.4256 | -85.3312
6.426 | -85.3312
6.4263 | -85.331
6.4266 | -85.3307
6.4269 | -85.3302
6.4272 | -85.3295
6.4275 | -85.3292
6.4278 | -85.3293
6.4281 | -85.3297
6.4284 | -85.3298
6.4288 | -85.33
6.4291 | -85.3302
6.4294 | -85.3302
6.4297 | -85.3306
6.43 | -85.331
6.4303 | -85.3312
6.4306 | -85.3317
6.4309 | -85.3318
6.4312 | -85.3319
6.4315 | -85.3322
6.4319 | -85.3325
6.4322 | -85.3327
6.4325 | -85.3326
6.4328 | -85.3323
6.4331 | -85.3341
6.4334 | -85.3343
6.4337 | -85.3346
6.434 | -85.3349
6.4343 | -85.3349
6.4347 | -85.3351
6.435 | -85.3352
6.4353 | -85.3341
6.4356 | -85.3334
6.4359 | -85.334
6.4362 | -85.3347
6.4365 | -85.335
6.4368 | -85.3353
6.4371 | -85.3356
6.4374 | -85.3359
6.4378 | -85.3363
6.4381 | -85.3368
6.4384 | -85.3376
6.4387 | -85.3382
6.439 | -85.3387
6.4393 | -85.3391
6.4396 | -85.34
6.4399 | -85.3416
6.4402 | -85.3431
6.4406 | -85.3442
6.4409 | -85.3454
6.4412 | -85.3468
6.4415 | -85.3478
6.4418 | -85.3499
6.4421 | -85.3502
6.4424 | -85.3509
6.4427 | -85.3515
6.443 | -85.3521
6.4433 | -85.3525
6.4437 | -85.3531
6.444 | -85.3537
6.4443 | -85.3547
6.4446 | -85.3556
6.4449 | -85.3563
6.4452 | -85.357
6.4455 | -85.3577
6.4458 | -85.3579
6.4461 | -85.3579
6.4465 | -85.3581
6.4468 | -85.3585
6.4471 | -85.3597
6.4474 | -85.3602
6.4477 | -85.3601
6.448 | -85.3607
6.4483 | -85.3617
6.4486 | -85.3629
6.4489 | -85.3641
6.4492 | -85.3654
6.4496 | -85.3665
6.4499 | -85.3668
6.4502 | -85.3676
6.4505 | -85.3685
6.4508 | -85.3707
6.4511 | -85.3717
6.4514 | -85.3725
6.4517 | -85.3733
6.452 | -85.3743
6.4524 | -85.3754
6.4527 | -85.3759
6.453 | -85.3767
6.4533 | -85.3773
6.4536 | -85.3779
6.4539 | -85.3791
6.4542 | -85.3798
6.4545 | -85.3812
6.4548 | -85.3823
6.4551 | -85.3836
6.4555 | -85.3852
6.4558 | -85.3866
6.4561 | -85.3876
6.4564 | -85.3881
6.4567 | -85.3884
6.457 | -85.3898
6.4573 | -85.391
6.4576 | -85.3926
6.4579 | -85.3938
6.4583 | -85.3946
6.4586 | -85.3958
6.4589 | -85.3968
6.4592 | -85.398
6.4595 | -85.3993
6.4598 | -85.402
6.4601 | -85.4034
6.4604 | -85.4048
6.4607 | -85.406
6.461 | -85.4065
6.4614 | -85.4076
6.4617 | -85.4095
6.462 | -85.4113
6.4623 | -85.4124
6.4626 | -85.4137
6.4629 | -85.415
6.4632 | -85.4155
6.4635 | -85.416
6.4638 | -85.4166
6.4642 | -85.4178
6.4645 | -85.4195
6.4648 | -85.4211
6.4651 | -85.4225
6.4654 | -85.4241
6.4657 | -85.4253
6.466 | -85.4267
6.4663 | -85.4281
6.4666 | -85.4293
6.4669 | -85.4308
6.4673 | -85.4323
6.4676 | -85.4339
6.4679 | -85.4353
6.4682 | -85.4368
6.4685 | -85.4382
6.4688 | -85.4408
6.4691 | -85.4423
6.4694 | -85.4437
6.4697 | -85.4452
6.4701 | -85.4467
6.4704 | -85.4483
6.4707 | -85.4495
6.471 | -85.451
6.4713 | -85.453
6.4716 | -85.4549
6.4719 | -85.456
6.4722 | -85.457
6.4725 | -85.4586
6.4729 | -85.4601
6.4732 | -85.4616
6.4735 | -85.4637
6.4738 | -85.465
6.4741 | -85.4663
6.4744 | -85.468
6.4747 | -85.4698
6.475 | -85.4716
6.4753 | -85.4737
6.4756 | -85.4754
6.476 | -85.477
6.4763 | -85.4788
6.4766 | -85.4808
6.4769 | -85.4827
6.4772 | -85.4844
6.4775 | -85.4862
6.4778 | -85.4893
6.4781 | -85.4915
6.4784 | -85.4939
6.4787 | -85.4958
6.4791 | -85.4975
6.4794 | -85.4993
6.4797 | -85.501
6.48 | -85.5027
6.4803 | -85.5046
6.4806 | -85.5064
6.4809 | -85.5085
6.4812 | -85.5103
6.4815 | -85.512
6.4819 | -85.5137
6.4822 | -85.5154
6.4825 | -85.5171
6.4828 | -85.519
6.4831 | -85.5208
6.4834 | -85.5224
6.4837 | -85.5244
6.484 | -85.5262
6.4843 | -85.5283
6.4847 | -85.5304
6.485 | -85.5321
6.4853 | -85.5342
6.4856 | -85.5363
6.4859 | -85.5382
6.4862 | -85.5401
6.4865 | -85.5421
6.4868 | -85.5453
6.4871 | -85.5475
6.4874 | -85.5499
6.4878 | -85.552
6.4881 | -85.5541
6.4884 | -85.5561
6.4887 | -85.5581
6.489 | -85.5601
6.4893 | -85.5621
6.4896 | -85.5642
6.4899 | -85.5663
6.4902 | -85.5683
6.4905 | -85.5703
6.4909 | -85.5724
6.4912 | -85.5747
6.4915 | -85.5766
6.4918 | -85.5783
6.4921 | -85.5802
6.4924 | -85.5822
6.4927 | -85.5846
6.493 | -85.5866
6.4933 | -85.5888
6.4937 | -85.591
6.494 | -85.5932
6.4943 | -85.5957
6.4946 | -85.5983
6.4949 | -85.6005
6.4952 | -85.6008
6.4955 | -85.6027
6.4958 | -85.6058
6.4961 | -85.6081
6.4965 | -85.6105
6.4968 | -85.6132
6.4971 | -85.6154
6.4974 | -85.6179
6.4977 | -85.6199
6.498 | -85.6223
6.4983 | -85.6244
6.4986 | -85.6263
6.4989 | -85.6282
6.4992 | -85.63
6.4996 | -85.6322
6.4999 | -85.6343
6.5002 | -85.636
6.5005 | -85.6384
6.5008 | -85.6407
6.5011 | -85.6428
6.5014 | -85.6452
6.5017 | -85.6476
6.502 | -85.6498
6.5024 | -85.652
6.5027 | -85.6543
6.503 | -85.6567
6.5033 | -85.6592
6.5036 | -85.6615
6.5039 | -85.664
6.5042 | -85.6665
6.5045 | -85.6694
6.5048 | -85.6715
6.5051 | -85.6739
6.5055 | -85.6764
6.5058 | -85.6787
6.5061 | -85.6807
6.5064 | -85.6831
6.5067 | -85.6852
6.507 | -85.6869
6.5073 | -85.6893
6.5076 | -85.6916
6.5079 | -85.694
6.5083 | -85.6966
6.5086 | -85.6988
6.5089 | -85.7014
6.5092 | -85.704
6.5095 | -85.7066
6.5098 | -85.7091
6.5101 | -85.7112
6.5104 | -85.7134
6.5107 | -85.7157
6.511 | -85.7183
6.5114 | -85.7211
6.5117 | -85.7236
6.512 | -85.7262
6.5123 | -85.7289
6.5126 | -85.7319
6.5129 | -85.7344
6.5132 | -85.7372
6.5135 | -85.7394
6.5138 | -85.7421
6.5142 | -85.7441
6.5145 | -85.7466
6.5148 | -85.7493
6.5151 | -85.752
6.5154 | -85.7549
6.5157 | -85.758
6.516 | -85.7606
6.5163 | -85.7631
6.5166 | -85.7662
6.5169 | -85.7694
6.5173 | -85.772
6.5176 | -85.7744
6.5179 | -85.7768
6.5182 | -85.7795
6.5185 | -85.7817
6.5188 | -85.7842
6.5191 | -85.787
6.5194 | -85.7897
6.5197 | -85.7923
6.5201 | -85.7942
6.5204 | -85.7964
6.5207 | -85.7991
6.521 | -85.8019
6.5213 | -85.8047
6.5216 | -85.8079
6.5219 | -85.8109
6.5222 | -85.8136
6.5225 | -85.817
6.5228 | -85.8192
6.5232 | -85.8211
6.5235 | -85.8238
6.5238 | -85.826
6.5241 | -85.8285
6.5244 | -85.832
6.5247 | -85.8347
6.525 | -85.8374
6.5253 | -85.8406
6.5256 | -85.8437
6.526 | -85.8463
6.5263 | -85.849
6.5266 | -85.8522
6.5269 | -85.8557
6.5272 | -85.8595
6.5275 | -85.8632
6.5278 | -85.8661
6.5281 | -85.8691
6.5284 | -85.8722
6.5287 | -85.8751
6.5291 | -85.8778
6.5294 | -85.88
6.5297 | -85.882
6.53 | -85.8849
6.5303 | -85.8875
6.5306 | -85.89
6.5309 | -85.8928
6.5312 | -85.8958
6.5315 | -85.8985
6.5319 | -85.9012
6.5322 | -85.9032
6.5325 | -85.9054
6.5328 | -85.908
6.5331 | -85.9099
6.5334 | -85.9118
6.5337 | -85.9145
6.534 | -85.9174
6.5343 | -85.9198
6.5346 | -85.9227
6.535 | -85.9257
6.5353 | -85.9287
6.5356 | -85.9317
6.5359 | -85.9347
6.5362 | -85.9389
6.5365 | -85.9426
6.5368 | -85.9455
6.5371 | -85.9483
6.5374 | -85.9512
6.5378 | -85.9543
6.5381 | -85.9572
6.5384 | -85.9597
6.5387 | -85.9625
6.539 | -85.9654
6.5393 | -85.9685
6.5396 | -85.9711
6.5399 | -85.9734
6.5402 | -85.9762
6.5405 | -85.9795
6.5409 | -85.9825
6.5412 | -85.9855
6.5415 | -85.9884
6.5418 | -85.9912
6.5421 | -85.994
6.5424 | -85.997
6.5427 | -85.9998
6.543 | -86.0027
6.5433 | -86.0056
6.5437 | -86.0086
6.544 | -86.0124
6.5443 | -86.0159
6.5446 | -86.0188
6.5449 | -86.0218
6.5452 | -86.0254
6.5455 | -86.0282
6.5458 | -86.032
6.5461 | -86.0347
6.5464 | -86.0372
6.5468 | -86.0405
6.5471 | -86.0436
6.5474 | -86.0464
6.5477 | -86.0493
6.548 | -86.0524
6.5483 | -86.0552
6.5486 | -86.0585
6.5489 | -86.0618
6.5492 | -86.065
6.5496 | -86.068
6.5499 | -86.0711
6.5502 | -86.0742
6.5505 | -86.0772
6.5508 | -86.0798
6.5511 | -86.0826
6.5514 | -86.0856
6.5517 | -86.0886
6.552 | -86.0913
6.5523 | -86.0944
6.5527 | -86.0973
6.553 | -86.1002
6.5533 | -86.1032
6.5536 | -86.1061
6.5539 | -86.1089
6.5542 | -86.1107
6.5545 | -86.1139
6.5548 | -86.1166
6.5551 | -86.119
6.5555 | -86.1218
6.5558 | -86.1244
6.5561 | -86.1272
6.5564 | -86.1297
6.5567 | -86.1322
6.557 | -86.1349
6.5573 | -86.1379
6.5576 | -86.1406
6.5579 | -86.144
6.5582 | -86.1471
6.5586 | -86.1504
6.5589 | -86.1534
6.5592 | -86.1566
6.5595 | -86.1592
6.5598 | -86.1623
6.5601 | -86.1653
6.5604 | -86.1679
6.5607 | -86.1711
6.561 | -86.1754
6.5614 | -86.1803
6.5617 | -86.1847
6.562 | -86.1882
6.5623 | -86.1924
6.5626 | -86.1962
6.5629 | -86.199
6.5632 | -86.201
6.5635 | -86.2047
6.5638 | -86.2098
6.5641 | -86.2131
6.5645 | -86.216
6.5648 | -86.2187
6.5651 | -86.2216
6.5654 | -86.2246
6.5657 | -86.2275
6.566 | -86.2304
6.5663 | -86.2335
6.5666 | -86.2363
6.5669 | -86.2392
6.5673 | -86.2418
6.5676 | -86.2449
6.5679 | -86.2478
6.5682 | -86.2508
6.5685 | -86.2537
6.5688 | -86.2567
6.5691 | -86.2597
6.5694 | -86.2634
6.5697 | -86.2682
6.57 | -86.2718
6.5704 | -86.2746
6.5707 | -86.2771
6.571 | -86.2796
6.5713 | -86.2826
6.5716 | -86.2856
6.5719 | -86.288
6.5722 | -86.2898
6.5725 | -86.2923
6.5728 | -86.2953
6.5732 | -86.2981
6.5735 | -86.3007
6.5738 | -86.3038
6.5741 | -86.3066
6.5744 | -86.3094
6.5747 | -86.3122
6.575 | -86.3152
6.5753 | -86.3184
6.5756 | -86.3205
6.5759 | -86.3225
6.5763 | -86.3247
6.5766 | -86.3271
6.5769 | -86.33
6.5772 | -86.3331
6.5775 | -86.3362
6.5778 | -86.339
6.5781 | -86.3417
6.5784 | -86.3444
6.5787 | -86.3467
6.5791 | -86.3494
6.5794 | -86.352
6.5797 | -86.3548
6.58 | -86.3581
6.5803 | -86.3611
6.5806 | -86.3639
6.5809 | -86.3672
6.5812 | -86.3711
6.5815 | -86.3744
6.5818 | -86.3775
6.5822 | -86.3802
6.5825 | -86.383
6.5828 | -86.3862
6.5831 | -86.3898
6.5834 | -86.3924
6.5837 | -86.395
6.584 | -86.398
6.5843 | -86.4013
6.5846 | -86.4041
6.585 | -86.4069
6.5853 | -86.4092
6.5856 | -86.4122
6.5859 | -86.4163
6.5862 | -86.4201
6.5865 | -86.4225
6.5868 | -86.4253
6.5871 | -86.4281
6.5874 | -86.4322
6.5877 | -86.435
6.5881 | -86.4367
6.5884 | -86.4394
6.5887 | -86.442
6.589 | -86.4448
6.5893 | -86.4475
6.5896 | -86.4509
6.5899 | -86.4543
6.5902 | -86.4576
6.5905 | -86.4603
6.5909 | -86.463
6.5912 | -86.4662
6.5915 | -86.4685
6.5918 | -86.4713
6.5921 | -86.474
6.5924 | -86.4777
6.5927 | -86.4817
6.593 | -86.4847
6.5933 | -86.4877
6.5936 | -86.4904
6.594 | -86.4932
6.5943 | -86.4955
6.5946 | -86.4978
6.5949 | -86.4991
6.5952 | -86.5019
6.5955 | -86.5048
6.5958 | -86.5076
6.5961 | -86.5102
6.5964 | -86.5129
6.5968 | -86.5156
6.5971 | -86.5181
6.5974 | -86.5213
6.5977 | -86.5243
6.598 | -86.5267
6.5983 | -86.5292
6.5986 | -86.5313
6.5989 | -86.5337
6.5992 | -86.5368
6.5995 | -86.5392
6.5999 | -86.541
6.6002 | -86.5427
6.6005 | -86.5452
6.6008 | -86.5473
6.6011 | -86.5496
6.6014 | -86.5525
6.6017 | -86.5556
6.602 | -86.5592
6.6023 | -86.5621
6.6027 | -86.5657
6.603 | -86.5692
6.6033 | -86.5714
6.6036 | -86.5735
6.6039 | -86.5758
6.6042 | -86.5786
6.6045 | -86.5817
6.6048 | -86.5844
6.6051 | -86.5869
6.6054 | -86.5893
6.6058 | -86.593
6.6061 | -86.5977
6.6064 | -86.6005
6.6067 | -86.6028
6.607 | -86.6056
6.6073 | -86.6085
6.6076 | -86.6113
6.6079 | -86.6141
6.6082 | -86.6165
6.6086 | -86.6188
6.6089 | -86.6211
6.6092 | -86.6234
6.6095 | -86.6259
6.6098 | -86.629
6.6101 | -86.6315
6.6104 | -86.6338
6.6107 | -86.6357
6.611 | -86.637
6.6113 | -86.6401
6.6117 | -86.6419
6.612 | -86.6432
6.6123 | -86.6448
6.6126 | -86.6478
6.6129 | -86.6504
6.6132 | -86.653
6.6135 | -86.6555
6.6138 | -86.6583
6.6141 | -86.6611
6.6145 | -86.6631
6.6148 | -86.665
6.6151 | -86.6674
6.6154 | -86.6716
6.6157 | -86.6753
6.616 | -86.6777
6.6163 | -86.6799
6.6166 | -86.6814
6.6169 | -86.6836
6.6173 | -86.6859
6.6176 | -86.6883
6.6179 | -86.6908
6.6182 | -86.6928
6.6185 | -86.6951
6.6188 | -86.6972
6.6191 | -86.6993
6.6194 | -86.7015
6.6197 | -86.7037
6.62 | -86.7063
6.6204 | -86.7096
6.6207 | -86.7115
6.621 | -86.7131
6.6213 | -86.7154
6.6216 | -86.7188
6.6219 | -86.7219
6.6222 | -86.7243
6.6225 | -86.7271
6.6228 | -86.7292
6.6231 | -86.7309
6.6235 | -86.7329
6.6238 | -86.7352
6.6241 | -86.7382
6.6244 | -86.7412
6.6247 | -86.7431
6.625 | -86.745
6.6253 | -86.7472
6.6256 | -86.7496
6.6259 | -86.7526
6.6263 | -86.7551
6.6266 | -86.7579
6.6269 | -86.7601
6.6272 | -86.7628
6.6275 | -86.7664
6.6278 | -86.7708
6.6281 | -86.7747
6.6284 | -86.7771
6.6287 | -86.7789
6.6291 | -86.7812
6.6294 | -86.7831
6.6297 | -86.785
6.63 | -86.7865
6.6303 | -86.7889
6.6306 | -86.7909
6.6309 | -86.7959
6.6312 | -86.7998
6.6315 | -86.8029
6.6318 | -86.8058
6.6322 | -86.8079
6.6325 | -86.8097
6.6328 | -86.8118
6.6331 | -86.8134
6.6334 | -86.8157
6.6337 | -86.8181
6.634 | -86.8202
6.6343 | -86.8217
6.6346 | -86.8254
6.6349 | -86.8301
6.6353 | -86.8326
6.6356 | -86.8357
6.6359 | -86.8393
6.6362 | -86.842
6.6365 | -86.8441
6.6368 | -86.8465
6.6371 | -86.8481
6.6374 | -86.8495
6.6377 | -86.8509
6.6381 | -86.8527
6.6384 | -86.8545
6.6387 | -86.8565
6.639 | -86.858
6.6393 | -86.8599
6.6396 | -86.8619
6.6399 | -86.8634
6.6402 | -86.8651
6.6405 | -86.8677
6.6409 | -86.8699
6.6412 | -86.8704
6.6415 | -86.8728
6.6418 | -86.8753
6.6421 | -86.8768
6.6424 | -86.8785
6.6427 | -86.8802
6.643 | -86.8823
6.6433 | -86.8852
6.6436 | -86.8871
6.644 | -86.8885
6.6443 | -86.8901
6.6446 | -86.8917
6.6449 | -86.8933
6.6452 | -86.8954
6.6455 | -86.8967
6.6458 | -86.8972
6.6461 | -86.8994
6.6464 | -86.9009
6.6468 | -86.9021
6.6471 | -86.9035
6.6474 | -86.9049
6.6477 | -86.9068
6.648 | -86.9082
6.6483 | -86.9099
6.6486 | -86.9115
6.6489 | -86.9128
6.6492 | -86.914
6.6495 | -86.9162
6.6499 | -86.9169
6.6502 | -86.9174
6.6505 | -86.9193
6.6508 | -86.922
6.6511 | -86.9244
6.6514 | -86.9245
6.6517 | -86.9258
6.652 | -86.927
6.6523 | -86.9282
6.6527 | -86.9295
6.653 | -86.932
6.6533 | -86.9337
6.6536 | -86.9359
6.6539 | -86.9379
6.6542 | -86.9394
6.6545 | -86.9408
6.6548 | -86.943
6.6551 | -86.9475
6.6554 | -86.9515
6.6558 | -86.9528
6.6561 | -86.9537
6.6564 | -86.9554
6.6567 | -86.9569
6.657 | -86.9578
6.6573 | -86.9591
6.6576 | -86.9601
6.6579 | -86.9615
6.6582 | -86.9628
6.6586 | -86.9639
6.6589 | -86.9652
6.6592 | -86.9668
6.6595 | -86.9687
6.6598 | -86.9703
6.6601 | -86.9715
6.6604 | -86.9727
6.6607 | -86.974
6.661 | -86.9748
6.6613 | -86.9767
6.6617 | -86.9788
6.662 | -86.9804
6.6623 | -86.9821
6.6626 | -86.9831
6.6629 | -86.9833
6.6632 | -86.9845
6.6635 | -86.9857
6.6638 | -86.9874
6.6641 | -86.9888
6.6645 | -86.9901
6.6648 | -86.9925
6.6651 | -86.995
6.6654 | -86.996
6.6657 | -86.9972
6.666 | -86.9983
6.6663 | -86.9996
6.6666 | -87.0021
6.6669 | -87.0043
6.6672 | -87.0055
6.6676 | -87.0067
6.6679 | -87.0079
6.6682 | -87.0093
6.6685 | -87.0106
6.6688 | -87.0117
6.6691 | -87.0129
6.6694 | -87.0143
6.6697 | -87.0154
6.67 | -87.0166
6.6704 | -87.0178
6.6707 | -87.0201
6.671 | -87.0217
6.6713 | -87.0229
6.6716 | -87.0241
6.6719 | -87.0263
6.6722 | -87.0274
6.6725 | -87.0286
6.6728 | -87.0294
6.6731 | -87.0313
6.6735 | -87.033
6.6738 | -87.0338
6.6741 | -87.0347
6.6744 | -87.0358
6.6747 | -87.0367
6.675 | -87.0377
6.6753 | -87.0402
6.6756 | -87.042
6.6759 | -87.0431
6.6763 | -87.0441
6.6766 | -87.0448
6.6769 | -87.0445
6.6772 | -87.0464
6.6775 | -87.0474
6.6778 | -87.0483
6.6781 | -87.0496
6.6784 | -87.0509
6.6787 | -87.0522
6.679 | -87.0534
6.6794 | -87.054
6.6797 | -87.0557
6.68 | -87.0577
6.6803 | -87.0585
6.6806 | -87.0602
6.6809 | -87.0613
6.6812 | -87.0616
6.6815 | -87.0622
6.6818 | -87.0648
6.6822 | -87.0664
6.6825 | -87.0674
6.6828 | -87.0689
6.6831 | -87.07
6.6834 | -87.0708
6.6837 | -87.0702
6.684 | -87.0728
6.6843 | -87.0762
6.6846 | -87.0782
6.6849 | -87.0809
6.6853 | -87.0833
6.6856 | -87.0844
6.6859 | -87.0854
6.6862 | -87.0864
6.6865 | -87.0876
6.6868 | -87.0885
6.6871 | -87.089
6.6874 | -87.0898
6.6877 | -87.0906
6.6881 | -87.0911
6.6884 | -87.0916
6.6887 | -87.0923
6.689 | -87.0932
6.6893 | -87.0945
6.6896 | -87.0961
6.6899 | -87.0976
6.6902 | -87.0985
6.6905 | -87.0992
6.6908 | -87.0999
6.6912 | -87.1008
6.6915 | -87.102
6.6918 | -87.1031
6.6921 | -87.1044
6.6924 | -87.1053
6.6927 | -87.1053
6.693 | -87.1057
6.6933 | -87.1062
6.6936 | -87.1066
6.694 | -87.1075
6.6943 | -87.1086
6.6946 | -87.1097
6.6949 | -87.1109
6.6952 | -87.1117
6.6955 | -87.1119
6.6958 | -87.1126
6.6961 | -87.1133
6.6964 | -87.1142
6.6967 | -87.1156
6.6971 | -87.1164
6.6974 | -87.1176
6.6977 | -87.1193
6.698 | -87.1213
6.6983 | -87.1226
6.6986 | -87.1226
6.6989 | -87.1228
6.6992 | -87.1239
6.6995 | -87.1243
6.6999 | -87.1248
6.7002 | -87.1257
6.7005 | -87.1266
6.7008 | -87.1275
6.7011 | -87.1284
6.7014 | -87.1291
6.7017 | -87.1288
6.702 | -87.1297
6.7023 | -87.1304
6.7026 | -87.1316
6.703 | -87.1324
6.7033 | -87.1331
6.7036 | -87.134
6.7039 | -87.1345
6.7042 | -87.1351
6.7045 | -87.136
6.7048 | -87.1366
6.7051 | -87.1371
6.7054 | -87.1374
6.7058 | -87.1376
6.7061 | -87.1384
6.7064 | -87.1388
6.7067 | -87.1391
6.707 | -87.1396
6.7073 | -87.14
6.7076 | -87.1404
6.7079 | -87.1418
6.7082 | -87.1444
6.7085 | -87.1469
6.7089 | -87.148
6.7092 | -87.1482
6.7095 | -87.1476
6.7098 | -87.1472
6.7101 | -87.1481
6.7104 | -87.1491
6.7107 | -87.1494
6.711 | -87.1503
6.7113 | -87.1509
6.7117 | -87.1514
6.712 | -87.152
6.7123 | -87.1526
6.7126 | -87.1531
6.7129 | -87.1542
6.7132 | -87.1557
6.7135 | -87.156
6.7138 | -87.1562
6.7141 | -87.1572
6.7144 | -87.1576
6.7148 | -87.158
6.7151 | -87.1583
6.7154 | -87.1587
6.7157 | -87.1593
6.716 | -87.1604
6.7163 | -87.1607
6.7166 | -87.1612
6.7169 | -87.1615
6.7172 | -87.1618
6.7176 | -87.1626
6.7179 | -87.1633
6.7182 | -87.1639
6.7185 | -87.1644
6.7188 | -87.1653
6.7191 | -87.166
6.7194 | -87.1672
6.7197 | -87.1678
6.72 | -87.1687
6.7203 | -87.1692
6.7207 | -87.1692
6.721 | -87.1694
6.7213 | -87.17
6.7216 | -87.1712
6.7219 | -87.1718
6.7222 | -87.1727
6.7225 | -87.1745
6.7228 | -87.1756
6.7231 | -87.176
6.7235 | -87.1762
6.7238 | -87.1766
6.7241 | -87.177
6.7244 | -87.1777
6.7247 | -87.1785
6.725 | -87.1791
6.7253 | -87.1796
6.7256 | -87.1806
6.7259 | -87.183
6.7262 | -87.1865
6.7266 | -87.1875
6.7269 | -87.188
6.7272 | -87.188
6.7275 | -87.188
6.7278 | -87.1884
6.7281 | -87.1888
6.7284 | -87.1892
6.7287 | -87.1894
6.729 | -87.1895
6.7294 | -87.1899
6.7297 | -87.1903
6.73 | -87.1905
6.7303 | -87.1914
6.7306 | -87.1925
6.7309 | -87.1927
6.7312 | -87.1927
6.7315 | -87.1939
6.7318 | -87.1955
6.7321 | -87.1962
6.7325 | -87.1969
6.7328 | -87.198
6.7331 | -87.1994
6.7334 | -87.2
6.7337 | -87.2004
6.734 | -87.2014
6.7343 | -87.2032
6.7346 | -87.2052
6.7349 | -87.2067
6.7353 | -87.2074
6.7356 | -87.2082
6.7359 | -87.2098
6.7362 | -87.2111
6.7365 | -87.2125
6.7368 | -87.2141
6.7371 | -87.2153
6.7374 | -87.2164
6.7377 | -87.2165
6.738 | -87.2164
6.7384 | -87.2167
6.7387 | -87.2175
6.739 | -87.2188
6.7393 | -87.219
6.7396 | -87.22
6.7399 | -87.2202
6.7402 | -87.2203
6.7405 | -87.2209
6.7408 | -87.2214
6.7412 | -87.2225
6.7415 | -87.2231
6.7418 | -87.2239
6.7421 | -87.2248
6.7424 | -87.2252
6.7427 | -87.2256
6.743 | -87.2257
6.7433 | -87.2255
6.7436 | -87.2259
6.7439 | -87.2262
6.7443 | -87.2264
6.7446 | -87.2267
6.7449 | -87.2267
6.7452 | -87.2264
6.7455 | -87.2264
6.7458 | -87.2269
6.7461 | -87.2269
6.7464 | -87.2267
6.7467 | -87.2269
6.7471 | -87.2276
6.7474 | -87.2284
6.7477 | -87.2286
6.748 | -87.2285
6.7483 | -87.2288
6.7486 | -87.2295
6.7489 | -87.2293
6.7492 | -87.2291
6.7495 | -87.2295
6.7498 | -87.2303
6.7502 | -87.2307
6.7505 | -87.2308
6.7508 | -87.2308
6.7511 | -87.2311
6.7514 | -87.2314
6.7517 | -87.2319
6.752 | -87.2326
6.7523 | -87.2328
6.7526 | -87.2327
6.753 | -87.2324
6.7533 | -87.2327
6.7536 | -87.2333
6.7539 | -87.2343
6.7542 | -87.235
6.7545 | -87.2367
6.7548 | -87.2382
6.7551 | -87.2384
6.7554 | -87.2393
6.7557 | -87.2402
6.7561 | -87.241
6.7564 | -87.2413
6.7567 | -87.2417
6.757 | -87.2422
6.7573 | -87.2429
6.7576 | -87.2438
6.7579 | -87.2447
6.7582 | -87.2452
6.7585 | -87.2454
6.7589 | -87.2454
6.7592 | -87.2454
6.7595 | -87.2455
6.7598 | -87.2457
6.7601 | -87.246
6.7604 | -87.2464
6.7607 | -87.2468
6.761 | -87.2468
6.7613 | -87.2469
6.7617 | -87.2469
6.762 | -87.2469
6.7623 | -87.248
6.7626 | -87.2497
6.7629 | -87.2503
6.7632 | -87.2507
6.7635 | -87.2516
6.7638 | -87.2518
6.7641 | -87.2521
6.7644 | -87.2531
6.7648 | -87.2537
6.7651 | -87.2549
6.7654 | -87.2556
6.7657 | -87.2561
6.766 | -87.2566
6.7663 | -87.2571
6.7666 | -87.2586
6.7669 | -87.2597
6.7672 | -87.26
6.7675 | -87.2603
6.7679 | -87.2608
6.7682 | -87.262
6.7685 | -87.2623
6.7688 | -87.2624
6.7691 | -87.2629
6.7694 | -87.2635
6.7697 | -87.2642
6.77 | -87.2653
6.7703 | -87.2666
6.7707 | -87.2679
6.771 | -87.2689
6.7713 | -87.2694
6.7716 | -87.2695
6.7719 | -87.2701
6.7722 | -87.2709
6.7725 | -87.2713
6.7728 | -87.2716
6.7731 | -87.2717
6.7735 | -87.2728
6.7738 | -87.2729
6.7741 | -87.2732
6.7744 | -87.2738
6.7747 | -87.2744
6.775 | -87.2757
6.7753 | -87.277
6.7756 | -87.2775
6.7759 | -87.2783
6.7762 | -87.2793
6.7766 | -87.2798
6.7769 | -87.2805
6.7772 | -87.2814
6.7775 | -87.2818
6.7778 | -87.2822
6.7781 | -87.2823
6.7784 | -87.2826
6.7787 | -87.2837
6.779 | -87.284
6.7793 | -87.2846
6.7797 | -87.2855
6.78 | -87.2862
6.7803 | -87.2862
6.7806 | -87.2867
6.7809 | -87.2871
6.7812 | -87.2875
6.7815 | -87.2873
6.7818 | -87.2868
6.7821 | -87.2868
6.7825 | -87.2877
6.7828 | -87.2892
6.7831 | -87.2904
6.7834 | -87.2908
6.7837 | -87.2911
6.784 | -87.2915
6.7843 | -87.2919
6.7846 | -87.2922
6.7849 | -87.2926
6.7853 | -87.2931
6.7856 | -87.2936
6.7859 | -87.294
6.7862 | -87.2941
6.7865 | -87.2939
6.7868 | -87.2949
6.7871 | -87.2956
6.7874 | -87.296
6.7877 | -87.2965
6.788 | -87.297
6.7884 | -87.2974
6.7887 | -87.2978
6.789 | -87.2979
6.7893 | -87.2984
6.7896 | -87.299
6.7899 | -87.2992
6.7902 | -87.2998
6.7905 | -87.3001
6.7908 | -87.3003
6.7912 | -87.3
6.7915 | -87.3004
6.7918 | -87.3009
6.7921 | -87.3013
6.7924 | -87.302
6.7927 | -87.3027
6.793 | -87.3033
6.7933 | -87.3038
6.7936 | -87.3046
6.7939 | -87.3051
6.7943 | -87.3054
6.7946 | -87.3063
6.7949 | -87.3074
6.7952 | -87.3077
6.7955 | -87.3076
6.7958 | -87.3079
6.7961 | -87.3083
6.7964 | -87.3087
6.7967 | -87.3099
6.7971 | -87.311
6.7974 | -87.3114
6.7977 | -87.3127
6.798 | -87.3138
6.7983 | -87.3139
6.7986 | -87.3148
6.7989 | -87.3167
6.7992 | -87.3174
6.7995 | -87.3181
6.7998 | -87.3191
6.8002 | -87.3198
6.8005 | -87.3208
6.8008 | -87.3215
6.8011 | -87.322
6.8014 | -87.3228
6.8017 | -87.3237
6.802 | -87.3246
6.8023 | -87.3253
6.8026 | -87.3262
6.803 | -87.327
6.8033 | -87.3275
6.8036 | -87.3279
6.8039 | -87.3286
6.8042 | -87.3291
6.8045 | -87.3304
6.8048 | -87.3316
6.8051 | -87.3324
6.8054 | -87.3336
6.8057 | -87.3341
6.8061 | -87.3348
6.8064 | -87.3355
6.8067 | -87.3369
6.807 | -87.3378
6.8073 | -87.3386
6.8076 | -87.3394
6.8079 | -87.3402
6.8082 | -87.3412
6.8085 | -87.3425
6.8089 | -87.3432
6.8092 | -87.3443
6.8095 | -87.3444
6.8098 | -87.3453
6.8101 | -87.3458
6.8104 | -87.3464
6.8107 | -87.3471
6.811 | -87.3477
6.8113 | -87.3489
6.8116 | -87.3498
6.812 | -87.3506
6.8123 | -87.3511
6.8126 | -87.3522
6.8129 | -87.353
6.8132 | -87.3537
6.8135 | -87.3545
6.8138 | -87.355
6.8141 | -87.3553
6.8144 | -87.3559
6.8148 | -87.3561
6.8151 | -87.3562
6.8154 | -87.3567
6.8157 | -87.3569
6.816 | -87.3571
6.8163 | -87.3587
6.8166 | -87.3614
6.8169 | -87.3627
6.8172 | -87.3634
6.8175 | -87.3641
6.8179 | -87.3646
6.8182 | -87.3658
6.8185 | -87.3663
6.8188 | -87.3675
6.8191 | -87.3684
6.8194 | -87.3688
6.8197 | -87.3697
6.82 | -87.3704
6.8203 | -87.3713
6.8207 | -87.3724
6.821 | -87.373
6.8213 | -87.3736
6.8216 | -87.3744
6.8219 | -87.3755
6.8222 | -87.3766
6.8225 | -87.3782
6.8228 | -87.3788
6.8231 | -87.3796
6.8234 | -87.3801
6.8238 | -87.3805
6.8241 | -87.3809
6.8244 | -87.3816
6.8247 | -87.3825
6.825 | -87.3838
6.8253 | -87.3848
6.8256 | -87.3857
6.8259 | -87.3867
6.8262 | -87.3877
6.8266 | -87.3882
6.8269 | -87.3882
6.8272 | -87.3898
6.8275 | -87.3909
6.8278 | -87.3918
6.8281 | -87.3927
6.8284 | -87.3936
6.8287 | -87.3945
6.829 | -87.3954
6.8293 | -87.3964
6.8297 | -87.3972
6.83 | -87.3979
6.8303 | -87.3988
6.8306 | -87.3997
6.8309 | -87.4007
6.8312 | -87.4013
6.8315 | -87.4019
6.8318 | -87.4029
6.8321 | -87.4038
6.8325 | -87.4047
6.8328 | -87.4054
6.8331 | -87.4063
6.8334 | -87.4075
6.8337 | -87.4087
6.834 | -87.4099
6.8343 | -87.411
6.8346 | -87.4119
6.8349 | -87.4128
6.8352 | -87.4139
6.8356 | -87.4147
6.8359 | -87.4158
6.8362 | -87.4167
6.8365 | -87.4178
6.8368 | -87.4193
6.8371 | -87.4206
6.8374 | -87.4218
6.8377 | -87.4232
6.838 | -87.4245
6.8384 | -87.4255
6.8387 | -87.4265
6.839 | -87.4274
6.8393 | -87.4283
6.8396 | -87.4295
6.8399 | -87.4309
6.8402 | -87.4319
6.8405 | -87.4332
6.8408 | -87.4349
6.8411 | -87.4361
6.8415 | -87.4371
6.8418 | -87.4383
6.8421 | -87.4395
6.8424 | -87.4401
6.8427 | -87.441
6.843 | -87.442
6.8433 | -87.443
6.8436 | -87.4441
6.8439 | -87.4452
6.8443 | -87.446
6.8446 | -87.4472
6.8449 | -87.4484
6.8452 | -87.4505
6.8455 | -87.4515
6.8458 | -87.4529
6.8461 | -87.4543
6.8464 | -87.4556
6.8467 | -87.4575
6.847 | -87.4592
6.8474 | -87.4605
6.8477 | -87.4613
6.848 | -87.4623
6.8483 | -87.4639
6.8486 | -87.465
6.8489 | -87.4661
6.8492 | -87.4683
6.8495 | -87.4703
6.8498 | -87.4717
6.8502 | -87.4739
6.8505 | -87.4759
6.8508 | -87.4778
6.8511 | -87.4795
6.8514 | -87.4805
6.8517 | -87.4818
6.852 | -87.4841
6.8523 | -87.4864
6.8526 | -87.4874
6.8529 | -87.4889
6.8533 | -87.4901
6.8536 | -87.4912
6.8539 | -87.4931
6.8542 | -87.4949
6.8545 | -87.4968
6.8548 | -87.4984
6.8551 | -87.5
6.8554 | -87.5016
6.8557 | -87.5027
6.8561 | -87.5038
6.8564 | -87.5048
6.8567 | -87.506
6.857 | -87.5075
6.8573 | -87.5087
6.8576 | -87.5096
6.8579 | -87.5106
6.8582 | -87.5121
6.8585 | -87.5128
6.8588 | -87.5136
6.8592 | -87.5151
6.8595 | -87.5168
6.8598 | -87.5183
6.8601 | -87.5198
6.8604 | -87.5212
6.8607 | -87.5229
6.861 | -87.5242
6.8613 | -87.5254
6.8616 | -87.5265
6.862 | -87.5281
6.8623 | -87.5297
6.8626 | -87.5311
6.8629 | -87.5332
6.8632 | -87.5344
6.8635 | -87.536
6.8638 | -87.5375
6.8641 | -87.5392
6.8644 | -87.5405
6.8647 | -87.5425
6.8651 | -87.5443
6.8654 | -87.5458
6.8657 | -87.5468
6.866 | -87.5477
6.8663 | -87.549
6.8666 | -87.5505
6.8669 | -87.5531
6.8672 | -87.5547
6.8675 | -87.5563
6.8679 | -87.5578
6.8682 | -87.5587
6.8685 | -87.5601
6.8688 | -87.5623
6.8691 | -87.5648
6.8694 | -87.5666
6.8697 | -87.5685
6.87 | -87.5701
6.8703 | -87.5726
6.8706 | -87.5754
6.871 | -87.5768
6.8713 | -87.5782
6.8716 | -87.5801
6.8719 | -87.5823
6.8722 | -87.5848
6.8725 | -87.5869
6.8728 | -87.5885
6.8731 | -87.5897
6.8734 | -87.591
6.8738 | -87.5923
6.8741 | -87.5942
6.8744 | -87.5961
6.8747 | -87.5975
6.875 | -87.5988
6.8753 | -87.6002
6.8756 | -87.6012
6.8759 | -87.6025
6.8762 | -87.6043
6.8765 | -87.6056
6.8769 | -87.6074
6.8772 | -87.6096
6.8775 | -87.611
6.8778 | -87.6126
6.8781 | -87.6141
6.8784 | -87.6156
6.8787 | -87.617
6.879 | -87.618
6.8793 | -87.6195
6.8797 | -87.6209
6.88 | -87.6224
6.8803 | -87.6242
6.8806 | -87.6261
6.8809 | -87.6278
6.8812 | -87.6283
6.8815 | -87.63
6.8818 | -87.6324
6.8821 | -87.6345
6.8824 | -87.6362
6.8828 | -87.6381
6.8831 | -87.6401
6.8834 | -87.6414
6.8837 | -87.6429
6.884 | -87.6445
6.8843 | -87.646
6.8846 | -87.6473
6.8849 | -87.6487
6.8852 | -87.6504
6.8856 | -87.6519
6.8859 | -87.6533
6.8862 | -87.6548
6.8865 | -87.6564
6.8868 | -87.6571
6.8871 | -87.6587
6.8874 | -87.6601
6.8877 | -87.6618
6.888 | -87.6635
6.8883 | -87.6652
6.8887 | -87.667
6.889 | -87.6696
6.8893 | -87.6719
6.8896 | -87.6735
6.8899 | -87.6758
6.8902 | -87.6775
6.8905 | -87.679
6.8908 | -87.6807
6.8911 | -87.6821
6.8915 | -87.6833
6.8918 | -87.685
6.8921 | -87.6868
6.8924 | -87.6885
6.8927 | -87.6899
6.893 | -87.6913
6.8933 | -87.6935
6.8936 | -87.6958
6.8939 | -87.6976
6.8942 | -87.6992
6.8946 | -87.7008
6.8949 | -87.7034
6.8952 | -87.7063
6.8955 | -87.7085
6.8958 | -87.71
6.8961 | -87.7113
6.8964 | -87.7131
6.8967 | -87.7153
6.897 | -87.7173
6.8974 | -87.719
6.8977 | -87.7213
6.898 | -87.7255
6.8983 | -87.7287
6.8986 | -87.7299
6.8989 | -87.7309
6.8992 | -87.7322
6.8995 | -87.7338
6.8998 | -87.7357
6.9002 | -87.7373
6.9005 | -87.7393
6.9008 | -87.7407
6.9011 | -87.742
6.9014 | -87.7435
6.9017 | -87.7441
6.902 | -87.7455
6.9023 | -87.7474
6.9026 | -87.7493
6.9029 | -87.751
6.9033 | -87.7531
6.9036 | -87.7556
6.9039 | -87.7572
6.9042 | -87.7588
6.9045 | -87.7608
6.9048 | -87.7629
6.9051 | -87.7647
6.9054 | -87.7669
6.9057 | -87.7688
6.906 | -87.7705
6.9064 | -87.7724
6.9067 | -87.7743
6.907 | -87.7763
6.9073 | -87.7778
6.9076 | -87.7793
6.9079 | -87.7811
6.9082 | -87.7828
6.9085 | -87.7844
6.9088 | -87.7855
6.9092 | -87.7875
6.9095 | -87.7902
6.9098 | -87.7927
6.9101 | -87.7955
6.9104 | -87.7981
6.9107 | -87.8002
6.911 | -87.8028
6.9113 | -87.8056
6.9116 | -87.8075
6.912 | -87.8092
6.9123 | -87.811
6.9126 | -87.8129
6.9129 | -87.8146
6.9132 | -87.8164
6.9135 | -87.8192
6.9138 | -87.8218
6.9141 | -87.8236
6.9144 | -87.8253
6.9147 | -87.8278
6.9151 | -87.8294
6.9154 | -87.8307
6.9157 | -87.8318
6.916 | -87.8337
6.9163 | -87.8359
6.9166 | -87.8383
6.9169 | -87.8409
6.9172 | -87.8427
6.9175 | -87.8448
6.9178 | -87.8473
6.9182 | -87.8495
6.9185 | -87.8516
6.9188 | -87.8532
6.9191 | -87.8555
6.9194 | -87.8573
6.9197 | -87.8593
6.92 | -87.861
6.9203 | -87.8626
6.9206 | -87.8653
6.921 | -87.8685
6.9213 | -87.8704
6.9216 | -87.872
6.9219 | -87.8734
6.9222 | -87.8749
6.9225 | -87.877
6.9228 | -87.8793
6.9231 | -87.8812
6.9234 | -87.8828
6.9238 | -87.8847
6.9241 | -87.8866
6.9244 | -87.8886
6.9247 | -87.8906
6.925 | -87.8925
6.9253 | -87.895
6.9256 | -87.8975
6.9259 | -87.9004
6.9262 | -87.9026
6.9265 | -87.9042
6.9269 | -87.9056
6.9272 | -87.9075
6.9275 | -87.9095
6.9278 | -87.912
6.9281 | -87.9151
6.9284 | -87.9182
6.9287 | -87.9212
6.929 | -87.9227
6.9293 | -87.9237
6.9297 | -87.9253
6.93 | -87.9273
6.9303 | -87.9294
6.9306 | -87.9322
6.9309 | -87.934
6.9312 | -87.9357
6.9315 | -87.9376
6.9318 | -87.9392
6.9321 | -87.9407
6.9324 | -87.9424
6.9328 | -87.9444
6.9331 | -87.9462
6.9334 | -87.9488
6.9337 | -87.9514
6.934 | -87.9534
6.9343 | -87.9555
6.9346 | -87.9578
6.9349 | -87.96
6.9352 | -87.9618
6.9356 | -87.964
6.9359 | -87.9657
6.9362 | -87.9675
6.9365 | -87.9701
6.9368 | -87.9721
6.9371 | -87.9743
6.9374 | -87.9771
6.9377 | -87.9798
6.938 | -87.9818
6.9383 | -87.9835
6.9387 | -87.9855
6.939 | -87.9876
6.9393 | -87.9899
6.9396 | -87.992
6.9399 | -87.9941
6.9402 | -87.996
6.9405 | -87.9978
6.9408 | -87.9994
6.9411 | -88.0013
6.9415 | -88.0035
6.9418 | -88.0054
6.9421 | -88.0075
6.9424 | -88.0095
6.9427 | -88.0116
6.943 | -88.0136
6.9433 | -88.0155
6.9436 | -88.0181
6.9439 | -88.0202
6.9442 | -88.0224
6.9446 | -88.0249
6.9449 | -88.0268
6.9452 | -88.0288
6.9455 | -88.031
6.9458 | -88.033
6.9461 | -88.0351
6.9464 | -88.0373
6.9467 | -88.0399
6.947 | -88.0426
6.9474 | -88.0448
6.9477 | -88.0469
6.948 | -88.0478
6.9483 | -88.0494
6.9486 | -88.0508
6.9489 | -88.0509
6.9492 | -88.0521
6.9495 | -88.0544
6.9498 | -88.0564
6.9501 | -88.0582
6.9505 | -88.0602
6.9508 | -88.0623
6.9511 | -88.064
6.9514 | -88.0656
6.9517 | -88.0681
6.952 | -88.0701
6.9523 | -88.0723
6.9526 | -88.0748
6.9529 | -88.0765
6.9533 | -88.0786
6.9536 | -88.0805
6.9539 | -88.0821
6.9542 | -88.0832
6.9545 | -88.0845
6.9548 | -88.0863
6.9551 | -88.088
6.9554 | -88.0899
6.9557 | -88.0917
6.956 | -88.0935
6.9564 | -88.0957
6.9567 | -88.0979
6.957 | -88.1001
6.9573 | -88.1022
6.9576 | -88.1043
6.9579 | -88.1062
6.9582 | -88.1084
6.9585 | -88.1113
6.9588 | -88.114
6.9592 | -88.1159
6.9595 | -88.1177
6.9598 | -88.1199
6.9601 | -88.1214
6.9604 | -88.123
6.9607 | -88.125
6.961 | -88.1273
6.9613 | -88.1293
6.9616 | -88.1315
6.9619 | -88.1342
6.9623 | -88.1395
6.9626 | -88.1415
6.9629 | -88.1425
6.9632 | -88.1445
6.9635 | -88.1467
6.9638 | -88.1489
6.9641 | -88.1511
6.9644 | -88.1532
6.9647 | -88.1553
6.9651 | -88.1579
6.9654 | -88.1604
6.9657 | -88.1629
6.966 | -88.1655
6.9663 | -88.1682
6.9666 | -88.1708
6.9669 | -88.1729
6.9672 | -88.1751
6.9675 | -88.1781
6.9678 | -88.1805
6.9682 | -88.1826
6.9685 | -88.185
6.9688 | -88.1872
6.9691 | -88.1891
6.9694 | -88.1914
6.9697 | -88.1938
6.97 | -88.1959
6.9703 | -88.1973
6.9706 | -88.1986
6.971 | -88.2005
6.9713 | -88.2027
6.9716 | -88.2045
6.9719 | -88.2063
6.9722 | -88.2084
6.9725 | -88.2114
6.9728 | -88.2137
6.9731 | -88.2151
6.9734 | -88.217
6.9737 | -88.2192
6.9741 | -88.2215
6.9744 | -88.2235
6.9747 | -88.2261
6.975 | -88.2291
6.9753 | -88.2305
6.9756 | -88.2315
6.9759 | -88.2336
6.9762 | -88.2357
6.9765 | -88.2382
6.9769 | -88.2401
6.9772 | -88.2417
6.9775 | -88.2436
6.9778 | -88.2451
6.9781 | -88.248
6.9784 | -88.2512
6.9787 | -88.2535
6.979 | -88.256
6.9793 | -88.2581
6.9796 | -88.2576
6.98 | -88.2596
6.9803 | -88.261
6.9806 | -88.2629
6.9809 | -88.2648
6.9812 | -88.2668
6.9815 | -88.2687
6.9818 | -88.2718
6.9821 | -88.2748
6.9824 | -88.2763
6.9828 | -88.2793
6.9831 | -88.2816
6.9834 | -88.2827
6.9837 | -88.2837
6.984 | -88.2857
6.9843 | -88.2876
6.9846 | -88.2896
6.9849 | -88.2916
6.9852 | -88.2923
6.9855 | -88.2941
6.9859 | -88.2962
6.9862 | -88.2989
6.9865 | -88.3002
6.9868 | -88.3016
6.9871 | -88.3043
6.9874 | -88.3064
6.9877 | -88.3083
6.988 | -88.3103
6.9883 | -88.3122
6.9887 | -88.3146
6.989 | -88.3185
6.9893 | -88.3208
6.9896 | -88.321
6.9899 | -88.3208
6.9902 | -88.3221
6.9905 | -88.3239
6.9908 | -88.3252
6.9911 | -88.3269
6.9914 | -88.329
6.9918 | -88.3303
6.9921 | -88.333
6.9924 | -88.3367
6.9927 | -88.3382
6.993 | -88.3398
6.9933 | -88.3418
6.9936 | -88.3433
6.9939 | -88.3455
6.9942 | -88.3483
6.9946 | -88.3499
6.9949 | -88.3501
6.9952 | -88.3514
6.9955 | -88.3533
6.9958 | -88.3548
6.9961 | -88.3565
6.9964 | -88.3586
6.9967 | -88.361
6.997 | -88.3627
6.9973 | -88.3646
6.9977 | -88.3664
6.998 | -88.3678
6.9983 | -88.369
6.9986 | -88.371
6.9989 | -88.3736
6.9992 | -88.3763
6.9995 | -88.3783
6.9998 | -88.3801
7.0001 | -88.3825
7.0005 | -88.3842
7.0008 | -88.3857
7.0011 | -88.3879
7.0014 | -88.3905
7.0017 | -88.3937
7.002 | -88.3973
7.0023 | -88.4
7.0026 | -88.4022
7.0029 | -88.4055
7.0032 | -88.4086
7.0036 | -88.4101
7.0039 | -88.4119
7.0042 | -88.4145
7.0045 | -88.4165
7.0048 | -88.4182
7.0051 | -88.4199
7.0054 | -88.4213
7.0057 | -88.4228
7.006 | -88.4244
7.0064 | -88.4262
7.0067 | -88.4279
7.007 | -88.4295
7.0073 | -88.4307
7.0076 | -88.433
7.0079 | -88.4351
7.0082 | -88.4369
7.0085 | -88.4394
7.0088 | -88.4413
7.0091 | -88.4427
7.0095 | -88.4444
7.0098 | -88.4466
7.0101 | -88.4491
7.0104 | -88.451
7.0107 | -88.4528
7.011 | -88.4544
7.0113 | -88.456
7.0116 | -88.4577
7.0119 | -88.4593
7.0123 | -88.461
7.0126 | -88.4625
7.0129 | -88.4638
7.0132 | -88.4654
7.0135 | -88.4677
7.0138 | -88.4697
7.0141 | -88.4715
7.0144 | -88.4731
7.0147 | -88.4751
7.015 | -88.4766
7.0154 | -88.4772
7.0157 | -88.479
7.016 | -88.4805
7.0163 | -88.482
7.0166 | -88.4831
7.0169 | -88.4852
7.0172 | -88.4876
7.0175 | -88.4894
7.0178 | -88.4912
7.0182 | -88.4929
7.0185 | -88.4938
7.0188 | -88.4957
7.0191 | -88.4976
7.0194 | -88.4983
7.0197 | -88.4999
7.02 | -88.5015
7.0203 | -88.5031
7.0206 | -88.5046
7.0209 | -88.506
7.0213 | -88.5073
7.0216 | -88.509
7.0219 | -88.5114
7.0222 | -88.5135
7.0225 | -88.5148
7.0228 | -88.5164
7.0231 | -88.5184
7.0234 | -88.5205
7.0237 | -88.5229
7.0241 | -88.5243
7.0244 | -88.5255
7.0247 | -88.5269
7.025 | -88.5286
7.0253 | -88.531
7.0256 | -88.5347
7.0259 | -88.538
7.0262 | -88.5408
7.0265 | -88.5431
7.0268 | -88.5446
7.0272 | -88.5467
7.0275 | -88.5484
7.0278 | -88.5504
7.0281 | -88.5524
7.0284 | -88.554
7.0287 | -88.5552
7.029 | -88.5567
7.0293 | -88.5585
7.0296 | -88.5599
7.03 | -88.5614
7.0303 | -88.5627
7.0306 | -88.5641
7.0309 | -88.5652
7.0312 | -88.5665
7.0315 | -88.5675
7.0318 | -88.5687
7.0321 | -88.5696
7.0324 | -88.5703
7.0327 | -88.5713
7.0331 | -88.5729
7.0334 | -88.574
7.0337 | -88.5754
7.034 | -88.5768
7.0343 | -88.5777
7.0346 | -88.5785
7.0349 | -88.5807
7.0352 | -88.5835
7.0355 | -88.5855
7.0359 | -88.587
7.0362 | -88.5883
7.0365 | -88.5896
7.0368 | -88.5911
7.0371 | -88.5926
7.0374 | -88.594
7.0377 | -88.5947
7.038 | -88.5958
7.0383 | -88.5972
7.0386 | -88.5987
7.039 | -88.5996
7.0393 | -88.6011
7.0396 | -88.6025
7.0399 | -88.6034
7.0402 | -88.6037
7.0405 | -88.6042
7.0408 | -88.6053
7.0411 | -88.6056
7.0414 | -88.6061
7.0418 | -88.6073
7.0421 | -88.6088
7.0424 | -88.6099
7.0427 | -88.6116
7.043 | -88.613
7.0433 | -88.6147
7.0436 | -88.6167
7.0439 | -88.6182
7.0442 | -88.6195
7.0446 | -88.6211
7.0449 | -88.6236
7.0452 | -88.6257
7.0455 | -88.6267
7.0458 | -88.6278
7.0461 | -88.63
7.0464 | -88.6314
7.0467 | -88.6326
7.047 | -88.6336
7.0473 | -88.6338
7.0477 | -88.6347
7.048 | -88.6358
7.0483 | -88.6382
7.0486 | -88.6407
7.0489 | -88.6419
7.0492 | -88.644
7.0495 | -88.6456
7.0498 | -88.646
7.0501 | -88.6468
7.0504 | -88.6479
7.0508 | -88.6492
7.0511 | -88.6501
7.0514 | -88.6518
7.0517 | -88.6532
7.052 | -88.6552
7.0523 | -88.6558
7.0526 | -88.6552
7.0529 | -88.6549
7.0532 | -88.6558
7.0536 | -88.657
7.0539 | -88.6582
7.0542 | -88.6594
7.0545 | -88.66
7.0548 | -88.6605
7.0551 | -88.6615
7.0554 | -88.6618
7.0557 | -88.6626
7.056 | -88.6636
7.0564 | -88.6637
7.0567 | -88.6649
7.057 | -88.6667
7.0573 | -88.6685
7.0576 | -88.6693
7.0579 | -88.6702
7.0582 | -88.671
7.0585 | -88.6721
7.0588 | -88.6728
7.0591 | -88.6735
7.0595 | -88.6745
7.0598 | -88.6756
7.0601 | -88.6764
7.0604 | -88.6778
7.0607 | -88.6784
7.061 | -88.6801
7.0613 | -88.6814
7.0616 | -88.6829
7.0619 | -88.6849
7.0622 | -88.6858
7.0626 | -88.6869
7.0629 | -88.6875
7.0632 | -88.6879
7.0635 | -88.6885
7.0638 | -88.6895
7.0641 | -88.6899
7.0644 | -88.6909
7.0647 | -88.6916
7.065 | -88.6923
7.0654 | -88.6937
7.0657 | -88.6952
7.066 | -88.6963
7.0663 | -88.6972
7.0666 | -88.6975
7.0669 | -88.698
7.0672 | -88.6981
7.0675 | -88.699
7.0678 | -88.7
7.0682 | -88.7012
7.0685 | -88.7021
7.0688 | -88.7036
7.0691 | -88.7035
7.0694 | -88.7058
7.0697 | -88.7065
7.07 | -88.7074
7.0703 | -88.7081
7.0706 | -88.7087
7.0709 | -88.7095
7.0713 | -88.7101
7.0716 | -88.7112
7.0719 | -88.7115
7.0722 | -88.7123
7.0725 | -88.7135
7.0728 | -88.7147
7.0731 | -88.7167
7.0734 | -88.7174
7.0737 | -88.7181
7.0741 | -88.7191
7.0744 | -88.7214
7.0747 | -88.7237
7.075 | -88.7247
7.0753 | -88.7258
7.0756 | -88.7259
7.0759 | -88.7268
7.0762 | -88.7279
7.0765 | -88.7288
7.0768 | -88.7293
7.0772 | -88.7296
7.0775 | -88.7301
7.0778 | -88.7308
7.0781 | -88.7308
7.0784 | -88.7315
7.0787 | -88.7328
7.079 | -88.7338
7.0793 | -88.7344
7.0796 | -88.7347
7.08 | -88.735
7.0803 | -88.7351
7.0806 | -88.7355
7.0809 | -88.7358
7.0812 | -88.7362
7.0815 | -88.7369
7.0818 | -88.7383
7.0821 | -88.7401
7.0824 | -88.7405
7.0827 | -88.7406
7.0831 | -88.7413
7.0834 | -88.7417
7.0837 | -88.7425
7.084 | -88.7431
7.0843 | -88.7435
7.0846 | -88.7441
7.0849 | -88.7447
7.0852 | -88.7455
7.0855 | -88.7467
7.0859 | -88.748
7.0862 | -88.7501
7.0865 | -88.7512
7.0868 | -88.7533
7.0871 | -88.7532
7.0874 | -88.7545
7.0877 | -88.7549
7.088 | -88.7554
7.0883 | -88.7548
7.0886 | -88.7554
7.089 | -88.7563
7.0893 | -88.7568
7.0896 | -88.7573
7.0899 | -88.7583
7.0902 | -88.7595
7.0905 | -88.7603
7.0908 | -88.7611
7.0911 | -88.7615
7.0914 | -88.7613
7.0918 | -88.7617
7.0921 | -88.7622
7.0924 | -88.7634
7.0927 | -88.7632
7.093 | -88.7635
7.0933 | -88.7642
7.0936 | -88.7641
7.0939 | -88.7644
7.0942 | -88.7646
7.0945 | -88.7651
7.0949 | -88.7651
7.0952 | -88.7651
7.0955 | -88.7648
7.0958 | -88.7646
7.0961 | -88.7647
7.0964 | -88.765
7.0967 | -88.7662
7.097 | -88.7668
7.0973 | -88.7671
7.0977 | -88.7671
7.098 | -88.7663
7.0983 | -88.7661
7.0986 | -88.7661
7.0989 | -88.7667
7.0992 | -88.7675
7.0995 | -88.7685
7.0998 | -88.7689
7.1001 | -88.7691
7.1004 | -88.7694
7.1008 | -88.7695
7.1011 | -88.7701
7.1014 | -88.771
7.1017 | -88.7714
7.102 | -88.7715
7.1023 | -88.7715
7.1026 | -88.7717
7.1029 | -88.7726
7.1032 | -88.7728
7.1036 | -88.7727
7.1039 | -88.7725
7.1042 | -88.7726
7.1045 | -88.7725
7.1048 | -88.7721
7.1051 | -88.772
7.1054 | -88.7723
7.1057 | -88.7723
7.106 | -88.7727
7.1063 | -88.7735
7.1067 | -88.7738
7.107 | -88.774
7.1073 | -88.7742
7.1076 | -88.7742
7.1079 | -88.7742
7.1082 | -88.7747
7.1085 | -88.7748
7.1088 | -88.7747
7.1091 | -88.7752
7.1095 | -88.7759
7.1098 | -88.7762
7.1101 | -88.7763
7.1104 | -88.7755
7.1107 | -88.7751
7.111 | -88.7758
7.1113 | -88.7765
7.1116 | -88.7762
7.1119 | -88.7764
7.1122 | -88.7764
7.1126 | -88.7769
7.1129 | -88.777
7.1132 | -88.7772
7.1135 | -88.7776
7.1138 | -88.7774
7.1141 | -88.7777
7.1144 | -88.7785
7.1147 | -88.7793
7.115 | -88.78
7.1154 | -88.78
7.1157 | -88.7801
7.116 | -88.7801
7.1163 | -88.78
7.1166 | -88.7803
7.1169 | -88.7809
7.1172 | -88.7806
7.1175 | -88.7803
7.1178 | -88.7806
7.1181 | -88.7808
7.1185 | -88.7812
7.1188 | -88.7829
7.1191 | -88.7833
7.1194 | -88.7835
7.1197 | -88.7837
7.12 | -88.7833
7.1203 | -88.7835
7.1206 | -88.7831
7.1209 | -88.7824
7.1213 | -88.7819
7.1216 | -88.7815
7.1219 | -88.7806
7.1222 | -88.7794
7.1225 | -88.7789
7.1228 | -88.778
7.1231 | -88.778
7.1234 | -88.7799
7.1237 | -88.7821
7.124 | -88.7825
7.1244 | -88.7829
7.1247 | -88.7835
7.125 | -88.7841
7.1253 | -88.7842
7.1256 | -88.7841
7.1259 | -88.7835
7.1262 | -88.7831
7.1265 | -88.7828
7.1268 | -88.7825
7.1272 | -88.7827
7.1275 | -88.7825
7.1278 | -88.7825
7.1281 | -88.7839
7.1284 | -88.7849
7.1287 | -88.7849
7.129 | -88.7848
7.1293 | -88.7856
7.1296 | -88.7867
7.1299 | -88.7868
7.1303 | -88.7868
7.1306 | -88.7869
7.1309 | -88.7868
7.1312 | -88.7866
7.1315 | -88.7864
7.1318 | -88.7859
7.1321 | -88.7856
7.1324 | -88.7854
7.1327 | -88.7848
7.1331 | -88.7841
7.1334 | -88.7847
7.1337 | -88.7852
7.134 | -88.7851
7.1343 | -88.7854
7.1346 | -88.7861
7.1349 | -88.7857
7.1352 | -88.7854
7.1355 | -88.7849
7.1358 | -88.7851
7.1362 | -88.7858
7.1365 | -88.7861
7.1368 | -88.7857
7.1371 | -88.7852
7.1374 | -88.7845
7.1377 | -88.7845
7.138 | -88.7847
7.1383 | -88.7844
7.1386 | -88.7847
7.139 | -88.7849
7.1393 | -88.7841
7.1396 | -88.7834
7.1399 | -88.784
7.1402 | -88.7841
7.1405 | -88.7833
7.1408 | -88.7832
7.1411 | -88.7834
7.1414 | -88.783
7.1417 | -88.7829
7.1421 | -88.783
7.1424 | -88.7821
7.1427 | -88.7816
7.143 | -88.7811
7.1433 | -88.7806
7.1436 | -88.7799
7.1439 | -88.7784
7.1442 | -88.7781
7.1445 | -88.7781
7.1449 | -88.7773
7.1452 | -88.7763
7.1455 | -88.7759
7.1458 | -88.7754
7.1461 | -88.7748
7.1464 | -88.7736
7.1467 | -88.7731
7.147 | -88.7726
7.1473 | -88.7721
7.1476 | -88.7719
7.148 | -88.7716
7.1483 | -88.7708
7.1486 | -88.7703
7.1489 | -88.77
7.1492 | -88.7698
7.1495 | -88.7696
7.1498 | -88.7693
7.1501 | -88.7699
7.1504 | -88.7701
7.1508 | -88.7699
7.1511 | -88.7698
7.1514 | -88.7696
7.1517 | -88.7699
7.152 | -88.77
7.1523 | -88.769
7.1526 | -88.7686
7.1529 | -88.7682
7.1532 | -88.768
7.1535 | -88.7686
7.1539 | -88.7684
7.1542 | -88.7679
7.1545 | -88.7673
7.1548 | -88.7664
7.1551 | -88.7658
7.1554 | -88.7653
7.1557 | -88.7651
7.156 | -88.7646
7.1563 | -88.7638
7.1567 | -88.7643
7.157 | -88.7661
7.1573 | -88.7665
7.1576 | -88.7672
7.1579 | -88.767
7.1582 | -88.7662
7.1585 | -88.7653
7.1588 | -88.7648
7.1591 | -88.7646
7.1594 | -88.7634
7.1598 | -88.7628
7.1601 | -88.7629
7.1604 | -88.7627
7.1607 | -88.7626
7.161 | -88.7617
7.1613 | -88.7615
7.1616 | -88.762
7.1619 | -88.7633
7.1622 | -88.7633
7.1626 | -88.7628
7.1629 | -88.7612
7.1632 | -88.7603
7.1635 | -88.7598
7.1638 | -88.7591
7.1641 | -88.7599
7.1644 | -88.7599
7.1647 | -88.7596
7.165 | -88.7591
7.1653 | -88.7586
7.1657 | -88.7583
7.166 | -88.7575
7.1663 | -88.7566
7.1666 | -88.7579
7.1669 | -88.759
7.1672 | -88.7589
7.1675 | -88.7588
7.1678 | -88.7585
7.1681 | -88.7586
7.1685 | -88.7584
7.1688 | -88.758
7.1691 | -88.7581
7.1694 | -88.7576
7.1697 | -88.7568
7.17 | -88.7564
7.1703 | -88.7559
7.1706 | -88.7554
7.1709 | -88.7546
7.1712 | -88.7542
7.1716 | -88.7537
7.1719 | -88.7533
7.1722 | -88.7531
7.1725 | -88.7528
7.1728 | -88.7524
7.1731 | -88.7516
7.1734 | -88.7509
7.1737 | -88.7506
7.174 | -88.7505
7.1744 | -88.7505
7.1747 | -88.7505
7.175 | -88.7499
7.1753 | -88.7492
7.1756 | -88.7488
7.1759 | -88.7489
7.1762 | -88.7483
7.1765 | -88.7475
7.1768 | -88.746
7.1771 | -88.7458
7.1775 | -88.7453
7.1778 | -88.745
7.1781 | -88.7455
7.1784 | -88.7455
7.1787 | -88.745
7.179 | -88.7444
7.1793 | -88.7445
7.1796 | -88.7444
7.1799 | -88.7444
7.1803 | -88.7447
7.1806 | -88.7445
7.1809 | -88.7443
7.1812 | -88.7446
7.1815 | -88.7443
7.1818 | -88.7442
7.1821 | -88.7443
7.1824 | -88.7436
7.1827 | -88.7429
7.183 | -88.7432
7.1834 | -88.743
7.1837 | -88.743
7.184 | -88.743
7.1843 | -88.743
7.1846 | -88.7432
7.1849 | -88.7429
7.1852 | -88.7422
7.1855 | -88.7423
7.1858 | -88.7417
7.1862 | -88.7416
7.1865 | -88.7412
7.1868 | -88.7401
7.1871 | -88.7392
7.1874 | -88.7387
7.1877 | -88.7388
7.188 | -88.739
7.1883 | -88.7389
7.1886 | -88.739
7.189 | -88.7389
7.1893 | -88.7383
7.1896 | -88.738
7.1899 | -88.7379
7.1902 | -88.7375
7.1905 | -88.7373
7.1908 | -88.7369
7.1911 | -88.7363
7.1914 | -88.736
7.1917 | -88.7363
7.1921 | -88.7354
7.1924 | -88.7347
7.1927 | -88.7344
7.193 | -88.734
7.1933 | -88.7339
7.1936 | -88.7333
7.1939 | -88.7322
7.1942 | -88.7301
7.1945 | -88.7272
7.1948 | -88.7268
7.1952 | -88.7268
7.1955 | -88.7262
7.1958 | -88.7257
7.1961 | -88.7252
7.1964 | -88.7252
7.1967 | -88.7257
7.197 | -88.7249
7.1973 | -88.7242
7.1976 | -88.7237
7.198 | -88.7236
7.1983 | -88.7236
7.1986 | -88.7232
7.1989 | -88.7228
7.1992 | -88.7223
7.1995 | -88.7217
7.1998 | -88.7212
7.2001 | -88.7209
7.2004 | -88.7201
7.2008 | -88.7198
7.2011 | -88.7192
7.2014 | -88.7187
7.2017 | -88.7184
7.202 | -88.7182
7.2023 | -88.7177
7.2026 | -88.7171
7.2029 | -88.7169
7.2032 | -88.7164
7.2035 | -88.7167
7.2039 | -88.7165
7.2042 | -88.7167
7.2045 | -88.7169
7.2048 | -88.7165
7.2051 | -88.7162
7.2054 | -88.7156
7.2057 | -88.715
7.206 | -88.7146
7.2063 | -88.7137
7.2066 | -88.7135
7.207 | -88.7135
7.2073 | -88.7134
7.2076 | -88.7131
7.2079 | -88.7127
7.2082 | -88.7123
7.2085 | -88.7124
7.2088 | -88.7124
7.2091 | -88.7116
7.2094 | -88.7114
7.2098 | -88.7106
7.2101 | -88.7101
7.2104 | -88.7108
7.2107 | -88.712
7.211 | -88.7126
7.2113 | -88.7121
7.2116 | -88.7116
7.2119 | -88.7116
7.2122 | -88.7116
7.2126 | -88.7117
7.2129 | -88.7116
7.2132 | -88.7114
7.2135 | -88.7106
7.2138 | -88.7092
7.2141 | -88.7086
7.2144 | -88.7085
7.2147 | -88.7082
7.215 | -88.7079
7.2153 | -88.7079
7.2157 | -88.7075
7.216 | -88.7068
7.2163 | -88.7065
7.2166 | -88.7058
7.2169 | -88.7038
7.2172 | -88.7016
7.2175 | -88.7012
7.2178 | -88.7009
7.2181 | -88.7006
7.2185 | -88.7001
7.2188 | -88.6999
7.2191 | -88.6997
7.2194 | -88.6988
7.2197 | -88.6983
7.22 | -88.698
7.2203 | -88.6976
7.2206 | -88.6972
7.2209 | -88.6972
7.2212 | -88.6973
7.2216 | -88.6967
7.2219 | -88.6967
7.2222 | -88.6969
7.2225 | -88.6968
7.2228 | -88.6973
7.2231 | -88.6968
7.2234 | -88.6963
7.2237 | -88.6961
7.224 | -88.6958
7.2244 | -88.6952
7.2247 | -88.6946
7.225 | -88.6946
7.2253 | -88.6942
7.2256 | -88.694
7.2259 | -88.6934
7.2262 | -88.6929
7.2265 | -88.6927
7.2268 | -88.6925
7.2271 | -88.6923
7.2275 | -88.6918
7.2278 | -88.6916
7.2281 | -88.6924
7.2284 | -88.692
7.2287 | -88.6916
7.229 | -88.6917
7.2293 | -88.6913
7.2296 | -88.6917
7.2299 | -88.6921
7.2303 | -88.6918
7.2306 | -88.6919
7.2309 | -88.6916
7.2312 | -88.6919
7.2315 | -88.6925
7.2318 | -88.6927
7.2321 | -88.6924
7.2324 | -88.6922
7.2327 | -88.6923
7.233 | -88.6923
7.2334 | -88.6923
7.2337 | -88.6928
7.234 | -88.6929
7.2343 | -88.6934
7.2346 | -88.6939
7.2349 | -88.6938
7.2352 | -88.6934
7.2355 | -88.693
7.2358 | -88.6928
7.2362 | -88.693
7.2365 | -88.6931
7.2368 | -88.6925
7.2371 | -88.6923
7.2374 | -88.6918
7.2377 | -88.6914
7.238 | -88.6917
7.2383 | -88.6924
7.2386 | -88.6919
7.2389 | -88.6906
7.2393 | -88.69
7.2396 | -88.6897
7.2399 | -88.6898
7.2402 | -88.6891
7.2405 | -88.6887
7.2408 | -88.6887
7.2411 | -88.6882
7.2414 | -88.6878
7.2417 | -88.6877
7.2421 | -88.6869
7.2424 | -88.6863
7.2427 | -88.6857
7.243 | -88.6855
7.2433 | -88.6856
7.2436 | -88.685
7.2439 | -88.6841
7.2442 | -88.6839
7.2445 | -88.6831
7.2448 | -88.6836
7.2452 | -88.6838
7.2455 | -88.6837
7.2458 | -88.6836
7.2461 | -88.6835
7.2464 | -88.6834
7.2467 | -88.6834
7.247 | -88.6831
7.2473 | -88.6831
7.2476 | -88.6832
7.248 | -88.6832
7.2483 | -88.6833
7.2486 | -88.6831
7.2489 | -88.6829
7.2492 | -88.683
7.2495 | -88.6832
7.2498 | -88.6829
7.2501 | -88.6831
7.2504 | -88.6828
7.2507 | -88.6822
7.2511 | -88.6818
7.2514 | -88.6815
7.2517 | -88.6813
7.252 | -88.681
7.2523 | -88.6804
7.2526 | -88.6803
7.2529 | -88.6802
7.2532 | -88.6798
7.2535 | -88.6799
7.2539 | -88.6795
7.2542 | -88.6788
7.2545 | -88.6785
7.2548 | -88.6786
7.2551 | -88.6786
7.2554 | -88.6781
7.2557 | -88.6776
7.256 | -88.6771
7.2563 | -88.677
7.2566 | -88.6766
7.257 | -88.6764
7.2573 | -88.6761
7.2576 | -88.676
7.2579 | -88.6762
7.2582 | -88.6769
7.2585 | -88.6765
7.2588 | -88.6749
7.2591 | -88.6733
7.2594 | -88.6732
7.2598 | -88.6732
7.2601 | -88.673
7.2604 | -88.6729
7.2607 | -88.6728
7.261 | -88.6728
7.2613 | -88.6727
7.2616 | -88.6726
7.2619 | -88.6726
7.2622 | -88.673
7.2625 | -88.6732
7.2629 | -88.673
7.2632 | -88.6729
7.2635 | -88.6729
7.2638 | -88.6728
7.2641 | -88.6729
7.2644 | -88.6731
7.2647 | -88.6733
7.265 | -88.6733
7.2653 | -88.6732
7.2657 | -88.6738
7.266 | -88.6738
7.2663 | -88.6745
7.2666 | -88.6745
7.2669 | -88.6744
7.2672 | -88.6744
7.2675 | -88.6749
7.2678 | -88.6764
7.2681 | -88.6773
7.2684 | -88.6772
7.2688 | -88.6771
7.2691 | -88.678
7.2694 | -88.6792
7.2697 | -88.6798
7.27 | -88.6797
7.2703 | -88.6798
7.2706 | -88.6799
7.2709 | -88.6798
7.2712 | -88.6797
7.2716 | -88.6797
7.2719 | -88.6788
7.2722 | -88.6781
7.2725 | -88.6781
7.2728 | -88.6782
7.2731 | -88.6783
7.2734 | -88.6781
7.2737 | -88.6778
7.274 | -88.6781
7.2743 | -88.6792
7.2747 | -88.6795
7.275 | -88.6793
7.2753 | -88.6799
7.2756 | -88.6799
7.2759 | -88.6799
7.2762 | -88.6797
7.2765 | -88.6798
7.2768 | -88.6798
7.2771 | -88.68
7.2775 | -88.6798
7.2778 | -88.6791
7.2781 | -88.6795
7.2784 | -88.68
7.2787 | -88.68
7.279 | -88.6804
7.2793 | -88.6814
7.2796 | -88.6822
7.2799 | -88.6824
7.2802 | -88.6822
7.2806 | -88.6821
7.2809 | -88.6819
7.2812 | -88.6818
7.2815 | -88.6817
7.2818 | -88.6819
7.2821 | -88.6821
7.2824 | -88.6822
7.2827 | -88.6824
7.283 | -88.6824
7.2834 | -88.6824
7.2837 | -88.6826
7.284 | -88.6833
7.2843 | -88.6841
7.2846 | -88.6847
7.2849 | -88.6853
7.2852 | -88.6856
7.2855 | -88.6856
7.2858 | -88.6854
7.2861 | -88.6853
7.2865 | -88.6848
7.2868 | -88.6848
7.2871 | -88.6855
7.2874 | -88.6858
7.2877 | -88.685
7.288 | -88.6845
7.2883 | -88.6843
7.2886 | -88.6838
7.2889 | -88.6834
7.2893 | -88.6837
7.2896 | -88.6841
7.2899 | -88.6841
7.2902 | -88.6841
7.2905 | -88.6842
7.2908 | -88.6842
7.2911 | -88.684
7.2914 | -88.6837
7.2917 | -88.6836
7.292 | -88.6835
7.2924 | -88.6833
7.2927 | -88.6836
7.293 | -88.6843
7.2933 | -88.6845
7.2936 | -88.6849
7.2939 | -88.6852
7.2942 | -88.6852
7.2945 | -88.6854
7.2948 | -88.6856
7.2952 | -88.6859
7.2955 | -88.6862
7.2958 | -88.6864
7.2961 | -88.6867
7.2964 | -88.6868
7.2967 | -88.6869
7.297 | -88.6873
7.2973 | -88.6876
7.2976 | -88.6878
7.2979 | -88.6879
7.2983 | -88.6879
7.2986 | -88.6879
7.2989 | -88.6882
7.2992 | -88.6886
7.2995 | -88.6884
7.2998 | -88.6888
7.3001 | -88.6894
7.3004 | -88.6897
7.3007 | -88.6889
7.3011 | -88.6879
7.3014 | -88.6885
7.3017 | -88.6892
7.302 | -88.6891
7.3023 | -88.6891
7.3026 | -88.6896
7.3029 | -88.6902
7.3032 | -88.6902
7.3035 | -88.6908
7.3038 | -88.6918
7.3042 | -88.6925
7.3045 | -88.6928
7.3048 | -88.6941
7.3051 | -88.6946
7.3054 | -88.6951
7.3057 | -88.6961
7.306 | -88.6974
7.3063 | -88.6983
7.3066 | -88.6984
7.307 | -88.6993
7.3073 | -88.7
7.3076 | -88.7
7.3079 | -88.7001
7.3082 | -88.7004
7.3085 | -88.7009
7.3088 | -88.7012
7.3091 | -88.7013
7.3094 | -88.6997
7.3097 | -88.6994
7.3101 | -88.6996
7.3104 | -88.7001
7.3107 | -88.6998
7.311 | -88.7004
7.3113 | -88.7007
7.3116 | -88.701
7.3119 | -88.7012
7.3122 | -88.7014
7.3125 | -88.7014
7.3129 | -88.7007
7.3132 | -88.7004
7.3135 | -88.7009
7.3138 | -88.7014
7.3141 | -88.7025
7.3144 | -88.7033
7.3147 | -88.704
7.315 | -88.7044
7.3153 | -88.7046
7.3156 | -88.7046
7.316 | -88.7048
7.3163 | -88.7051
7.3166 | -88.7052
7.3169 | -88.7053
7.3172 | -88.7058
7.3175 | -88.707
7.3178 | -88.7075
7.3181 | -88.7076
7.3184 | -88.7075
7.3188 | -88.7076
7.3191 | -88.7071
7.3194 | -88.707
7.3197 | -88.7072
7.32 | -88.708
7.3203 | -88.7083
7.3206 | -88.7087
7.3209 | -88.7094
7.3212 | -88.7095
7.3215 | -88.7094
7.3219 | -88.7098
7.3222 | -88.7101
7.3225 | -88.7108
7.3228 | -88.7113
7.3231 | -88.7113
7.3234 | -88.7119
7.3237 | -88.712
7.324 | -88.7125
7.3243 | -88.713
7.3247 | -88.7132
7.325 | -88.7134
7.3253 | -88.7135
7.3256 | -88.7132
7.3259 | -88.7133
7.3262 | -88.7138
7.3265 | -88.7141
7.3268 | -88.7146
7.3271 | -88.7146
7.3274 | -88.715
7.3278 | -88.7164
7.3281 | -88.7167
7.3284 | -88.7168
7.3287 | -88.7168
7.329 | -88.7175
7.3293 | -88.7177
7.3296 | -88.7179
7.3299 | -88.7181
7.3302 | -88.7182
7.3306 | -88.7185
7.3309 | -88.719
7.3312 | -88.7193
7.3315 | -88.7196
7.3318 | -88.7201
7.3321 | -88.7203
7.3324 | -88.7207
7.3327 | -88.721
7.333 | -88.7213
7.3334 | -88.7218
7.3337 | -88.7222
7.334 | -88.7224
7.3343 | -88.7231
7.3346 | -88.7236
7.3349 | -88.7241
7.3352 | -88.7245
7.3355 | -88.7249
7.3358 | -88.7251
7.3361 | -88.7255
7.3365 | -88.7257
7.3368 | -88.7253
7.3371 | -88.7254
7.3374 | -88.7259
7.3377 | -88.7262
7.338 | -88.7261
7.3383 | -88.726
7.3386 | -88.7264
7.3389 | -88.7271
7.3392 | -88.7275
7.3396 | -88.728
7.3399 | -88.7285
7.3402 | -88.7294
7.3405 | -88.7304
7.3408 | -88.7311
7.3411 | -88.7316
7.3414 | -88.7318
7.3417 | -88.732
7.342 | -88.7325
7.3424 | -88.733
7.3427 | -88.7331
7.343 | -88.7336
7.3433 | -88.734
7.3436 | -88.7349
7.3439 | -88.7358
7.3442 | -88.7361
7.3445 | -88.7365
7.3448 | -88.737
7.3452 | -88.7376
7.3455 | -88.738
7.3458 | -88.7385
7.3461 | -88.7392
7.3464 | -88.7399
7.3467 | -88.7406
7.347 | -88.7414
7.3473 | -88.7421
7.3476 | -88.7427
7.3479 | -88.7435
7.3483 | -88.7441
7.3486 | -88.7446
7.3489 | -88.7455
7.3492 | -88.7469
7.3495 | -88.7477
7.3498 | -88.7483
7.3501 | -88.7487
7.3504 | -88.7489
7.3507 | -88.7499
7.3511 | -88.7509
7.3514 | -88.7509
7.3517 | -88.7513
7.352 | -88.7521
7.3523 | -88.7525
7.3526 | -88.7532
7.3529 | -88.7539
7.3532 | -88.754
7.3535 | -88.7547
7.3538 | -88.7553
7.3542 | -88.7557
7.3545 | -88.7561
7.3548 | -88.757
7.3551 | -88.7577
7.3554 | -88.7584
7.3557 | -88.7593
7.356 | -88.7599
7.3563 | -88.7602
7.3566 | -88.76
7.357 | -88.7605
7.3573 | -88.7612
7.3576 | -88.7617
7.3579 | -88.7619
7.3582 | -88.7614
7.3585 | -88.762
7.3588 | -88.7618
7.3591 | -88.7612
7.3594 | -88.7618
7.3597 | -88.7625
7.3601 | -88.7624
7.3604 | -88.7628
7.3607 | -88.7633
7.361 | -88.7635
7.3613 | -88.7638
7.3616 | -88.7639
7.3619 | -88.7651
7.3622 | -88.7658
7.3625 | -88.7663
7.3629 | -88.7666
7.3632 | -88.7668
7.3635 | -88.7676
7.3638 | -88.7683
7.3641 | -88.7691
7.3644 | -88.7696
7.3647 | -88.7702
7.365 | -88.7706
7.3653 | -88.7714
7.3656 | -88.7723
7.366 | -88.7729
7.3663 | -88.7736
7.3666 | -88.7744
7.3669 | -88.7751
7.3672 | -88.7756
7.3675 | -88.7761
7.3678 | -88.7765
7.3681 | -88.7768
7.3684 | -88.7774
7.3688 | -88.778
7.3691 | -88.7788
7.3694 | -88.7796
7.3697 | -88.78
7.37 | -88.7806
7.3703 | -88.7815
7.3706 | -88.7821
7.3709 | -88.7828
7.3712 | -88.7828
7.3715 | -88.7832
7.3719 | -88.7841
7.3722 | -88.7848
7.3725 | -88.7855
7.3728 | -88.7862
7.3731 | -88.787
7.3734 | -88.7881
7.3737 | -88.7889
7.374 | -88.7895
7.3743 | -88.7899
7.3747 | -88.7897
7.375 | -88.7898
7.3753 | -88.7903
7.3756 | -88.7909
7.3759 | -88.7927
7.3762 | -88.7946
7.3765 | -88.7952
7.3768 | -88.7953
7.3771 | -88.7956
7.3774 | -88.796
7.3778 | -88.7966
7.3781 | -88.797
7.3784 | -88.7978
7.3787 | -88.7985
7.379 | -88.7988
7.3793 | -88.8002
7.3796 | -88.8012
7.3799 | -88.8022
7.3802 | -88.8024
7.3806 | -88.8031
7.3809 | -88.804
7.3812 | -88.8042
7.3815 | -88.805
7.3818 | -88.8068
7.3821 | -88.8076
7.3824 | -88.8087
7.3827 | -88.8101
7.383 | -88.8107
7.3833 | -88.8106
7.3837 | -88.8112
7.384 | -88.8122
7.3843 | -88.8132
7.3846 | -88.8147
7.3849 | -88.8158
7.3852 | -88.8166
7.3855 | -88.817
7.3858 | -88.817
7.3861 | -88.8176
7.3865 | -88.818
7.3868 | -88.8188
7.3871 | -88.8203
7.3874 | -88.8215
7.3877 | -88.8223
7.388 | -88.8228
7.3883 | -88.8234
7.3886 | -88.8243
7.3889 | -88.8247
7.3892 | -88.8252
7.3896 | -88.8258
7.3899 | -88.8265
7.3902 | -88.8274
7.3905 | -88.8281
7.3908 | -88.8288
7.3911 | -88.8296
7.3914 | -88.8306
7.3917 | -88.8319
7.392 | -88.8331
7.3924 | -88.834
7.3927 | -88.8343
7.393 | -88.8352
7.3933 | -88.836
7.3936 | -88.8366
7.3939 | -88.8375
7.3942 | -88.8386
7.3945 | -88.8401
7.3948 | -88.8418
7.3951 | -88.8433
7.3955 | -88.8444
7.3958 | -88.8453
7.3961 | -88.846
7.3964 | -88.8467
7.3967 | -88.8469
7.397 | -88.8477
7.3973 | -88.8487
7.3976 | -88.8493
7.3979 | -88.851
7.3983 | -88.8522
7.3986 | -88.8524
7.3989 | -88.853
7.3992 | -88.8538
7.3995 | -88.8548
7.3998 | -88.8555
7.4001 | -88.8564
7.4004 | -88.8572
7.4007 | -88.8584
7.401 | -88.8591
7.4014 | -88.8598
7.4017 | -88.8607
7.402 | -88.8612
7.4023 | -88.8616
7.4026 | -88.8622
7.4029 | -88.8631
7.4032 | -88.8641
7.4035 | -88.8648
7.4038 | -88.8655
7.4042 | -88.8659
7.4045 | -88.8668
7.4048 | -88.8684
7.4051 | -88.8695
7.4054 | -88.8708
7.4057 | -88.8724
7.406 | -88.8732
7.4063 | -88.8733
7.4066 | -88.8739
7.4069 | -88.8749
7.4073 | -88.8756
7.4076 | -88.8761
7.4079 | -88.8765
7.4082 | -88.8778
7.4085 | -88.8793
7.4088 | -88.8803
7.4091 | -88.881
7.4094 | -88.8823
7.4097 | -88.8831
7.4101 | -88.8841
7.4104 | -88.8848
7.4107 | -88.8855
7.411 | -88.8863
7.4113 | -88.887
7.4116 | -88.8878
7.4119 | -88.8884
7.4122 | -88.8897
7.4125 | -88.8911
7.4128 | -88.8926
7.4132 | -88.8936
7.4135 | -88.8938
7.4138 | -88.8932
7.4141 | -88.8935
7.4144 | -88.8943
7.4147 | -88.8954
7.415 | -88.8958
7.4153 | -88.8963
7.4156 | -88.8968
7.416 | -88.8977
7.4163 | -88.8985
7.4166 | -88.8993
7.4169 | -88.9001
7.4172 | -88.9007
7.4175 | -88.901
7.4178 | -88.9014
7.4181 | -88.902
7.4184 | -88.9017
7.4187 | -88.9013
7.4191 | -88.9023
7.4194 | -88.9034
7.4197 | -88.9044
7.42 | -88.9052
7.4203 | -88.906
7.4206 | -88.9064
7.4209 | -88.9073
7.4212 | -88.9084
7.4215 | -88.9086
7.4219 | -88.9093
7.4222 | -88.91
7.4225 | -88.9106
7.4228 | -88.9113
7.4231 | -88.9121
7.4234 | -88.9122
7.4237 | -88.9125
7.424 | -88.9133
7.4243 | -88.9133
7.4246 | -88.9141
7.425 | -88.9156
7.4253 | -88.9167
7.4256 | -88.9175
7.4259 | -88.9185
7.4262 | -88.9192
7.4265 | -88.9199
7.4268 | -88.92
7.4271 | -88.9215
7.4274 | -88.9229
7.4278 | -88.9247
7.4281 | -88.9257
7.4284 | -88.9265
7.4287 | -88.9275
7.429 | -88.9287
7.4293 | -88.9297
7.4296 | -88.9309
7.4299 | -88.932
7.4302 | -88.9335
7.4305 | -88.9348
7.4309 | -88.9353
7.4312 | -88.9359
7.4315 | -88.9369
7.4318 | -88.9375
7.4321 | -88.9382
7.4324 | -88.939
7.4327 | -88.9397
7.433 | -88.9404
7.4333 | -88.9412
7.4337 | -88.942
7.434 | -88.9431
7.4343 | -88.9438
7.4346 | -88.9449
7.4349 | -88.9463
7.4352 | -88.9477
7.4355 | -88.9485
7.4358 | -88.9492
7.4361 | -88.9499
7.4364 | -88.95
7.4368 | -88.9501
7.4371 | -88.9507
7.4374 | -88.9519
7.4377 | -88.9524
7.438 | -88.9522
7.4383 | -88.9524
7.4386 | -88.9532
7.4389 | -88.9544
7.4392 | -88.9552
7.4396 | -88.956
7.4399 | -88.9568
7.4402 | -88.9575
7.4405 | -88.9583
7.4408 | -88.959
7.4411 | -88.9594
7.4414 | -88.9598
7.4417 | -88.9604
7.442 | -88.9613
7.4423 | -88.9624
7.4427 | -88.9632
7.443 | -88.9641
7.4433 | -88.9649
7.4436 | -88.9658
7.4439 | -88.9667
7.4442 | -88.9678
7.4445 | -88.9686
7.4448 | -88.9694
7.4451 | -88.9704
7.4455 | -88.9715
7.4458 | -88.9724
7.4461 | -88.9734
7.4464 | -88.9742
7.4467 | -88.9751
7.447 | -88.976
7.4473 | -88.977
7.4476 | -88.9793
7.4479 | -88.9806
7.4482 | -88.9817
7.4486 | -88.9845
7.4489 | -88.9864
7.4492 | -88.9872
7.4495 | -88.9879
7.4498 | -88.9889
7.4501 | -88.9901
7.4504 | -88.9908
7.4507 | -88.9925
7.451 | -88.9932
7.4514 | -88.9934
7.4517 | -88.9944
7.452 | -88.9955
7.4523 | -88.9963
7.4526 | -88.9967
7.4529 | -88.9979
7.4532 | -88.9988
7.4535 | -88.9999
7.4538 | -89.001
7.4541 | -89.0029
7.4545 | -89.004
7.4548 | -89.0057
7.4551 | -89.0071
7.4554 | -89.0084
7.4557 | -89.0095
7.456 | -89.0108
7.4563 | -89.0119
7.4566 | -89.0128
7.4569 | -89.0137
7.4573 | -89.0145
7.4576 | -89.015
7.4579 | -89.0149
7.4582 | -89.016
7.4585 | -89.0171
7.4588 | -89.0178
7.4591 | -89.0186
7.4594 | -89.0199
7.4597 | -89.0212
7.46 | -89.0221
7.4604 | -89.0233
7.4607 | -89.0242
7.461 | -89.0245
7.4613 | -89.0246
7.4616 | -89.0253
7.4619 | -89.026
7.4622 | -89.0269
7.4625 | -89.0278
7.4628 | -89.0285
7.4632 | -89.0299
7.4635 | -89.0306
7.4638 | -89.0326
7.4641 | -89.0338
7.4644 | -89.0347
7.4647 | -89.0356
7.465 | -89.0363
7.4653 | -89.0371
7.4656 | -89.038
7.466 | -89.0388
7.4663 | -89.0394
7.4666 | -89.0402
7.4669 | -89.0411
7.4672 | -89.0416
7.4675 | -89.0419
7.4678 | -89.043
7.4681 | -89.0442
7.4684 | -89.0449
7.4687 | -89.0446
7.4691 | -89.0445
7.4694 | -89.0451
7.4697 | -89.0459
7.47 | -89.0468
7.4703 | -89.0478
7.4706 | -89.0486
7.4709 | -89.0495
7.4712 | -89.0512
7.4715 | -89.0529
7.4718 | -89.054
7.4722 | -89.0553
7.4725 | -89.0563
7.4728 | -89.0572
7.4731 | -89.0582
7.4734 | -89.0587
7.4737 | -89.0595
7.474 | -89.0598
7.4743 | -89.0608
7.4746 | -89.0622
7.475 | -89.0632
7.4753 | -89.0642
7.4756 | -89.0652
7.4759 | -89.0661
7.4762 | -89.0666
7.4765 | -89.0678
7.4768 | -89.0698
7.4771 | -89.0721
7.4774 | -89.0735
7.4778 | -89.0747
7.4781 | -89.0756
7.4784 | -89.076
7.4787 | -89.077
7.479 | -89.0781
7.4793 | -89.0797
7.4796 | -89.08
7.4799 | -89.0791
7.4802 | -89.0795
7.4805 | -89.0803
7.4809 | -89.081
7.4812 | -89.0814
7.4815 | -89.0842
7.4818 | -89.0859
7.4821 | -89.0865
7.4824 | -89.088
7.4827 | -89.0889
7.483 | -89.0898
7.4833 | -89.0907
7.4836 | -89.0916
7.484 | -89.0921
7.4843 | -89.0929
7.4846 | -89.0943
7.4849 | -89.0952
7.4852 | -89.0959
7.4855 | -89.0977
7.4858 | -89.0992
7.4861 | -89.0998
7.4864 | -89.1006
7.4868 | -89.101
7.4871 | -89.1014
7.4874 | -89.1029
7.4877 | -89.1035
7.488 | -89.1045
7.4883 | -89.1049
7.4886 | -89.1061
7.4889 | -89.1076
7.4892 | -89.108
7.4896 | -89.1091
7.4899 | -89.1103
7.4902 | -89.1115
7.4905 | -89.1126
7.4908 | -89.1136
7.4911 | -89.1144
7.4914 | -89.1153
7.4917 | -89.1161
7.492 | -89.118
7.4923 | -89.121
7.4927 | -89.1241
7.493 | -89.125
7.4933 | -89.1257
7.4936 | -89.1271
7.4939 | -89.1281
7.4942 | -89.1285
7.4945 | -89.1296
7.4948 | -89.131
7.4951 | -89.1317
7.4955 | -89.1331
7.4958 | -89.1343
7.4961 | -89.135
7.4964 | -89.1361
7.4967 | -89.1372
7.497 | -89.1378
7.4973 | -89.1388
7.4976 | -89.1413
7.4979 | -89.1429
7.4982 | -89.1435
7.4986 | -89.1444
7.4989 | -89.1451
7.4992 | -89.1441
7.4995 | -89.1445
7.4998 | -89.1452
7.5001 | -89.1457
7.5004 | -89.1461
7.5007 | -89.1468
7.501 | -89.1468
7.5014 | -89.1472
7.5017 | -89.1484
7.502 | -89.149
7.5023 | -89.15
7.5026 | -89.1516
7.5029 | -89.152
7.5032 | -89.1525
7.5035 | -89.1528
7.5038 | -89.1533
7.5041 | -89.1543
7.5045 | -89.1551
7.5048 | -89.1558
7.5051 | -89.1564
7.5054 | -89.1564
7.5057 | -89.1559
7.506 | -89.1561
7.5063 | -89.1568
7.5066 | -89.1576
7.5069 | -89.158
7.5073 | -89.1587
7.5076 | -89.1592
7.5079 | -89.1597
7.5082 | -89.1607
7.5085 | -89.1615
7.5088 | -89.1622
7.5091 | -89.1632
7.5094 | -89.1641
7.5097 | -89.1645
7.51 | -89.1651
7.5104 | -89.1655
7.5107 | -89.1662
7.511 | -89.1656
7.5113 | -89.1653
7.5116 | -89.1657
7.5119 | -89.1664
7.5122 | -89.167
7.5125 | -89.1676
7.5128 | -89.1684
7.5132 | -89.1693
7.5135 | -89.169
7.5138 | -89.1681
7.5141 | -89.1679
7.5144 | -89.1683
7.5147 | -89.1691
7.515 | -89.1697
7.5153 | -89.17
7.5156 | -89.1706
7.5159 | -89.1715
7.5163 | -89.1723
7.5166 | -89.1734
7.5169 | -89.1746
7.5172 | -89.1749
7.5175 | -89.1753
7.5178 | -89.1757
7.5181 | -89.1763
7.5184 | -89.1771
7.5187 | -89.178
7.5191 | -89.1779
7.5194 | -89.1786
7.5197 | -89.1795
7.52 | -89.1802
7.5203 | -89.1806
7.5206 | -89.1816
7.5209 | -89.1826
7.5212 | -89.1836
7.5215 | -89.1847
7.5218 | -89.1854
7.5222 | -89.1861
7.5225 | -89.1867
7.5228 | -89.1875
7.5231 | -89.1887
7.5234 | -89.1896
7.5237 | -89.1905
7.524 | -89.1919
7.5243 | -89.1923
7.5246 | -89.1928
7.525 | -89.1937
7.5253 | -89.1944
7.5256 | -89.195
7.5259 | -89.1957
7.5262 | -89.1965
7.5265 | -89.1972
7.5268 | -89.1977
7.5271 | -89.1974
7.5274 | -89.1981
7.5277 | -89.1987
7.5281 | -89.199
7.5284 | -89.1994
7.5287 | -89.2003
7.529 | -89.2015
7.5293 | -89.2018
7.5296 | -89.2021
7.5299 | -89.2025
7.5302 | -89.2034
7.5305 | -89.2038
7.5309 | -89.2044
7.5312 | -89.2049
7.5315 | -89.2055
7.5318 | -89.2061
7.5321 | -89.207
7.5324 | -89.2077
7.5327 | -89.208
7.533 | -89.2082
7.5333 | -89.2087
7.5336 | -89.2095
7.534 | -89.2103
7.5343 | -89.2111
7.5346 | -89.2116
7.5349 | -89.2119
7.5352 | -89.2123
7.5355 | -89.2129
7.5358 | -89.2138
7.5361 | -89.2144
7.5364 | -89.2154
7.5368 | -89.2169
7.5371 | -89.2182
7.5374 | -89.219
7.5377 | -89.2204
7.538 | -89.2218
7.5383 | -89.2225
7.5386 | -89.2233
7.5389 | -89.224
7.5392 | -89.2247
7.5395 | -89.2254
7.5399 | -89.2259
7.5402 | -89.227
7.5405 | -89.2281
7.5408 | -89.2284
7.5411 | -89.2287
7.5414 | -89.2291
7.5417 | -89.2298
7.542 | -89.2302
7.5423 | -89.2302
7.5427 | -89.2309
7.543 | -89.2314
7.5433 | -89.232
7.5436 | -89.2332
7.5439 | -89.234
7.5442 | -89.2345
7.5445 | -89.2349
7.5448 | -89.2355
7.5451 | -89.2361
7.5454 | -89.2369
7.5458 | -89.2376
7.5461 | -89.2387
7.5464 | -89.2404
7.5467 | -89.2419
7.547 | -89.2429
7.5473 | -89.2446
7.5476 | -89.2458
7.5479 | -89.2464
7.5482 | -89.2472
7.5486 | -89.2477
7.5489 | -89.2481
7.5492 | -89.2487
7.5495 | -89.2492
7.5498 | -89.2501
7.5501 | -89.2509
7.5504 | -89.2517
7.5507 | -89.2524
7.551 | -89.2537
7.5513 | -89.2549
7.5517 | -89.2557
7.552 | -89.2564
7.5523 | -89.2574
7.5526 | -89.2582
7.5529 | -89.2589
7.5532 | -89.2591
7.5535 | -89.2597
7.5538 | -89.2607
7.5541 | -89.2623
7.5545 | -89.2628
7.5548 | -89.2629
7.5551 | -89.2634
7.5554 | -89.2637
7.5557 | -89.2639
7.556 | -89.264
7.5563 | -89.2644
7.5566 | -89.2657
7.5569 | -89.2665
7.5572 | -89.2671
7.5576 | -89.2677
7.5579 | -89.2684
7.5582 | -89.269
7.5585 | -89.2697
7.5588 | -89.2703
7.5591 | -89.2708
7.5594 | -89.2713
7.5597 | -89.2739
7.56 | -89.2768
7.5604 | -89.278
7.5607 | -89.2783
7.561 | -89.2788
7.5613 | -89.2798
7.5616 | -89.2806
7.5619 | -89.2805
7.5622 | -89.2809
7.5625 | -89.2812
7.5628 | -89.2822
7.5631 | -89.2837
7.5635 | -89.2846
7.5638 | -89.286
7.5641 | -89.2865
7.5644 | -89.2864
7.5647 | -89.2866
7.565 | -89.2871
7.5653 | -89.2878
7.5656 | -89.2885
7.5659 | -89.2891
7.5663 | -89.2903
7.5666 | -89.291
7.5669 | -89.2914
7.5672 | -89.2916
7.5675 | -89.2919
7.5678 | -89.2923
7.5681 | -89.2929
7.5684 | -89.2937
7.5687 | -89.2941
7.569 | -89.2946
7.5694 | -89.2955
7.5697 | -89.2966
7.57 | -89.2972
7.5703 | -89.2975
7.5706 | -89.2981
7.5709 | -89.2988
7.5712 | -89.2997
7.5715 | -89.3003
7.5718 | -89.3008
7.5722 | -89.3019
7.5725 | -89.3026
7.5728 | -89.3032
7.5731 | -89.3036
7.5734 | -89.3037
7.5737 | -89.3036
7.574 | -89.3041
7.5743 | -89.3045
7.5746 | -89.3055
7.5749 | -89.3058
7.5753 | -89.3072
7.5756 | -89.3084
7.5759 | -89.3089
7.5762 | -89.3094
7.5765 | -89.3098
7.5768 | -89.3101
7.5771 | -89.31
7.5774 | -89.311
7.5777 | -89.3121
7.5781 | -89.3127
7.5784 | -89.3131
7.5787 | -89.3133
7.579 | -89.3135
7.5793 | -89.3139
7.5796 | -89.3149
7.5799 | -89.3156
7.5802 | -89.3169
7.5805 | -89.3183
7.5808 | -89.3189
7.5812 | -89.3195
7.5815 | -89.3203
7.5818 | -89.3206
7.5821 | -89.3208
7.5824 | -89.3209
7.5827 | -89.3212
7.583 | -89.3216
7.5833 | -89.3223
7.5836 | -89.3229
7.584 | -89.324
7.5843 | -89.3258
7.5846 | -89.3265
7.5849 | -89.3268
7.5852 | -89.3274
7.5855 | -89.3276
7.5858 | -89.328
7.5861 | -89.3285
7.5864 | -89.329
7.5867 | -89.3296
7.5871 | -89.3303
7.5874 | -89.331
7.5877 | -89.3315
7.588 | -89.3319
7.5883 | -89.3323
7.5886 | -89.3329
7.5889 | -89.3339
7.5892 | -89.3349
7.5895 | -89.3359
7.5899 | -89.3371
7.5902 | -89.3383
7.5905 | -89.3387
7.5908 | -89.3392
7.5911 | -89.34
7.5914 | -89.3406
7.5917 | -89.3411
7.592 | -89.3425
7.5923 | -89.3429
7.5926 | -89.3433
7.593 | -89.3429
7.5933 | -89.3427
7.5936 | -89.3434
7.5939 | -89.3436
7.5942 | -89.3439
7.5945 | -89.345
7.5948 | -89.3463
7.5951 | -89.3469
7.5954 | -89.3477
7.5958 | -89.3481
7.5961 | -89.3483
7.5964 | -89.3484
7.5967 | -89.3477
7.597 | -89.3487
7.5973 | -89.349
7.5976 | -89.3497
7.5979 | -89.3508
7.5982 | -89.3519
7.5985 | -89.3523
7.5989 | -89.3523
7.5992 | -89.3529
7.5995 | -89.3536
7.5998 | -89.3538
7.6001 | -89.3544
7.6004 | -89.355
7.6007 | -89.3555
7.601 | -89.3558
7.6013 | -89.3564
7.6017 | -89.3566
7.602 | -89.3569
7.6023 | -89.3574
7.6026 | -89.3585
7.6029 | -89.3591
7.6032 | -89.3596
7.6035 | -89.3603
7.6038 | -89.3607
7.6041 | -89.3613
7.6044 | -89.3618
7.6048 | -89.3626
7.6051 | -89.3636
7.6054 | -89.3646
7.6057 | -89.3657
7.606 | -89.3664
7.6063 | -89.3668
7.6066 | -89.3677
7.6069 | -89.3681
7.6072 | -89.3686
7.6076 | -89.3691
7.6079 | -89.3694
7.6082 | -89.3704
7.6085 | -89.372
7.6088 | -89.3728
7.6091 | -89.3736
7.6094 | -89.3743
7.6097 | -89.3749
7.61 | -89.3758
7.6104 | -89.3763
7.6107 | -89.3764
7.611 | -89.3767
7.6113 | -89.3773
7.6116 | -89.3779
7.6119 | -89.378
7.6122 | -89.3778
7.6125 | -89.3782
7.6128 | -89.3789
7.6131 | -89.3794
7.6135 | -89.3799
7.6138 | -89.3801
7.6141 | -89.3807
7.6144 | -89.3814
7.6147 | -89.3817
7.615 | -89.3818
7.6153 | -89.382
7.6156 | -89.383
7.6159 | -89.384
7.6162 | -89.3853
7.6166 | -89.3864
7.6169 | -89.3871
7.6172 | -89.3877
7.6175 | -89.3882
7.6178 | -89.3888
7.6181 | -89.3894
7.6184 | -89.3895
7.6187 | -89.3903
7.619 | -89.3912
7.6194 | -89.3919
7.6197 | -89.3925
7.62 | -89.3933
7.6203 | -89.3937
7.6206 | -89.3944
7.6209 | -89.3952
7.6212 | -89.3965
7.6215 | -89.3982
7.6218 | -89.3993
7.6222 | -89.3998
7.6225 | -89.4001
7.6228 | -89.4005
7.6231 | -89.4004
7.6234 | -89.4008
7.6237 | -89.4008
7.624 | -89.3998
7.6243 | -89.3995
7.6246 | -89.4004
7.6249 | -89.4009
7.6253 | -89.4014
7.6256 | -89.4013
7.6259 | -89.401
7.6262 | -89.4006
7.6265 | -89.401
7.6268 | -89.4015
7.6271 | -89.4019
7.6274 | -89.4023
7.6277 | -89.4031
7.628 | -89.4034
7.6284 | -89.4039
7.6287 | -89.4043
7.629 | -89.4051
7.6293 | -89.4062
7.6296 | -89.4067
7.6299 | -89.4069
7.6302 | -89.4066
7.6305 | -89.4069
7.6308 | -89.4075
7.6312 | -89.4079
7.6315 | -89.4079
7.6318 | -89.4083
7.6321 | -89.4087
7.6324 | -89.4095
7.6327 | -89.4107
7.633 | -89.4112
7.6333 | -89.4109
7.6336 | -89.4111
7.6339 | -89.4123
7.6343 | -89.4133
7.6346 | -89.4143
7.6349 | -89.4154
7.6352 | -89.416
7.6355 | -89.4163
7.6358 | -89.4168
7.6361 | -89.4172
7.6364 | -89.4176
7.6367 | -89.4178
7.6371 | -89.4182
7.6374 | -89.4187
7.6377 | -89.4193
7.638 | -89.4196
7.6383 | -89.4206
7.6386 | -89.4206
7.6389 | -89.4203
7.6392 | -89.4202
7.6395 | -89.4202
7.6399 | -89.4214
7.6402 | -89.4224
7.6405 | -89.423
7.6408 | -89.4235
7.6411 | -89.4238
7.6414 | -89.4241
7.6417 | -89.4249
7.642 | -89.4265
7.6423 | -89.4272
7.6426 | -89.4283
7.643 | -89.4291
7.6433 | -89.4298
7.6436 | -89.4307
7.6439 | -89.4312
7.6442 | -89.4316
7.6445 | -89.4325
7.6448 | -89.4325
7.6451 | -89.4324
7.6454 | -89.433
7.6458 | -89.4336
7.6461 | -89.4348
7.6464 | -89.4353
7.6467 | -89.4358
7.647 | -89.4364
7.6473 | -89.4371
7.6476 | -89.4376
7.6479 | -89.4381
7.6482 | -89.4388
7.6485 | -89.4395
7.6489 | -89.4401
7.6492 | -89.4404
7.6495 | -89.4407
7.6498 | -89.4418
7.6501 | -89.4431
7.6504 | -89.4443
7.6507 | -89.4445
7.651 | -89.4449
7.6513 | -89.4456
7.6517 | -89.4459
7.652 | -89.446
7.6523 | -89.4463
7.6526 | -89.4474
7.6529 | -89.4483
7.6532 | -89.4488
7.6535 | -89.4496
7.6538 | -89.4508
7.6541 | -89.4512
7.6544 | -89.4518
7.6548 | -89.4527
7.6551 | -89.4536
7.6554 | -89.4542
7.6557 | -89.4561
7.656 | -89.4574
7.6563 | -89.4576
7.6566 | -89.4573
7.6569 | -89.458
7.6572 | -89.4592
7.6576 | -89.4611
7.6579 | -89.4621
7.6582 | -89.4629
7.6585 | -89.4642
7.6588 | -89.4654
7.6591 | -89.4662
7.6594 | -89.4671
7.6597 | -89.4684
7.66 | -89.4695
7.6603 | -89.4707
7.6607 | -89.4714
7.661 | -89.4721
7.6613 | -89.4729
7.6616 | -89.4733
7.6619 | -89.4738
7.6622 | -89.475
7.6625 | -89.4758
7.6628 | -89.4762
7.6631 | -89.4768
7.6635 | -89.4773
7.6638 | -89.4778
7.6641 | -89.4787
7.6644 | -89.4793
7.6647 | -89.4798
7.665 | -89.4804
7.6653 | -89.4811
7.6656 | -89.4817
7.6659 | -89.4822
7.6662 | -89.4828
7.6666 | -89.4841
7.6669 | -89.4848
7.6672 | -89.4851
7.6675 | -89.4856
7.6678 | -89.4858
7.6681 | -89.4865
7.6684 | -89.4873
7.6687 | -89.4886
7.669 | -89.4896
7.6694 | -89.4907
7.6697 | -89.4917
7.67 | -89.4922
7.6703 | -89.4916
7.6706 | -89.4911
7.6709 | -89.4918
7.6712 | -89.4923
7.6715 | -89.4927
7.6718 | -89.4931
7.6721 | -89.4938
7.6725 | -89.4944
7.6728 | -89.4954
7.6731 | -89.4966
7.6734 | -89.4976
7.6737 | -89.4984
7.674 | -89.4986
7.6743 | -89.5001
7.6746 | -89.5016
7.6749 | -89.5022
7.6753 | -89.5028
7.6756 | -89.5036
7.6759 | -89.5043
7.6762 | -89.5045
7.6765 | -89.5051
7.6768 | -89.5058
7.6771 | -89.5064
7.6774 | -89.5071
7.6777 | -89.5091
7.678 | -89.5102
7.6784 | -89.5104
7.6787 | -89.5118
7.679 | -89.5131
7.6793 | -89.5138
7.6796 | -89.5146
7.6799 | -89.5152
7.6802 | -89.5159
7.6805 | -89.5167
7.6808 | -89.5173
7.6812 | -89.5184
7.6815 | -89.5204
7.6818 | -89.521
7.6821 | -89.5212
7.6824 | -89.5217
7.6827 | -89.5224
7.683 | -89.5227
7.6833 | -89.5229
7.6836 | -89.5236
7.6839 | -89.524
7.6843 | -89.5245
7.6846 | -89.5251
7.6849 | -89.5264
7.6852 | -89.5278
7.6855 | -89.5284
7.6858 | -89.529
7.6861 | -89.5301
7.6864 | -89.5315
7.6867 | -89.5319
7.6871 | -89.5325
7.6874 | -89.5335
7.6877 | -89.5341
7.688 | -89.5344
7.6883 | -89.5346
7.6886 | -89.5357
7.6889 | -89.5367
7.6892 | -89.5377
7.6895 | -89.5389
7.6898 | -89.5398
7.6902 | -89.5405
7.6905 | -89.5411
7.6908 | -89.5423
7.6911 | -89.5433
7.6914 | -89.5441
7.6917 | -89.5445
7.692 | -89.5456
7.6923 | -89.5462
7.6926 | -89.5477
7.693 | -89.5491
7.6933 | -89.5493
7.6936 | -89.5501
7.6939 | -89.5508
7.6942 | -89.5513
7.6945 | -89.5515
7.6948 | -89.5518
7.6951 | -89.5526
7.6954 | -89.5531
7.6957 | -89.5533
7.6961 | -89.5531
7.6964 | -89.554
7.6967 | -89.5542
7.697 | -89.5552
7.6973 | -89.5562
7.6976 | -89.5569
7.6979 | -89.5582
7.6982 | -89.5594
7.6985 | -89.5603
7.6989 | -89.5613
7.6992 | -89.5621
7.6995 | -89.5628
7.6998 | -89.5638
7.7001 | -89.565
7.7004 | -89.5654
7.7007 | -89.5661
7.701 | -89.5668
7.7013 | -89.5678
7.7016 | -89.5692
7.702 | -89.57
7.7023 | -89.5705
7.7026 | -89.5713
7.7029 | -89.5721
7.7032 | -89.5729
7.7035 | -89.5742
7.7038 | -89.5754
7.7041 | -89.5765
7.7044 | -89.5772
7.7048 | -89.578
7.7051 | -89.5788
7.7054 | -89.5796
7.7057 | -89.58
7.706 | -89.5806
7.7063 | -89.5816
7.7066 | -89.5821
7.7069 | -89.5828
7.7072 | -89.5836
7.7075 | -89.5841
7.7079 | -89.5846
7.7082 | -89.5854
7.7085 | -89.5864
7.7088 | -89.5881
7.7091 | -89.5909
7.7094 | -89.5924
7.7097 | -89.593
7.71 | -89.5938
7.7103 | -89.5946
7.7107 | -89.5954
7.711 | -89.5969
7.7113 | -89.5996
7.7116 | -89.6031
7.7119 | -89.6045
7.7122 | -89.6055
7.7125 | -89.6062
7.7128 | -89.6072
7.7131 | -89.6081
7.7134 | -89.6091
7.7138 | -89.6106
7.7141 | -89.6117
7.7144 | -89.6125
7.7147 | -89.6136
7.715 | -89.6149
7.7153 | -89.6155
7.7156 | -89.6157
7.7159 | -89.6167
7.7162 | -89.6172
7.7166 | -89.6177
7.7169 | -89.6186
7.7172 | -89.6196
7.7175 | -89.6208
7.7178 | -89.6217
7.7181 | -89.6229
7.7184 | -89.6239
7.7187 | -89.6249
7.719 | -89.6259
7.7193 | -89.6267
7.7197 | -89.6276
7.72 | -89.6283
7.7203 | -89.6295
7.7206 | -89.6302
7.7209 | -89.631
7.7212 | -89.6319
7.7215 | -89.6327
7.7218 | -89.6336
7.7221 | -89.6343
7.7225 | -89.6349
7.7228 | -89.6359
7.7231 | -89.6371
7.7234 | -89.638
7.7237 | -89.6389
7.724 | -89.6399
7.7243 | -89.641
7.7246 | -89.6422
7.7249 | -89.6428
7.7252 | -89.6437
7.7256 | -89.6446
7.7259 | -89.6453
7.7262 | -89.6463
7.7265 | -89.6472
7.7268 | -89.6475
7.7271 | -89.6484
7.7274 | -89.65
7.7277 | -89.651
7.728 | -89.652
7.7284 | -89.6529
7.7287 | -89.6532
7.729 | -89.6538
7.7293 | -89.6547
7.7296 | -89.656
7.7299 | -89.6576
7.7302 | -89.6587
7.7305 | -89.6598
7.7308 | -89.6609
7.7311 | -89.6612
7.7315 | -89.6614
7.7318 | -89.662
7.7321 | -89.663
7.7324 | -89.6643
7.7327 | -89.6663
7.733 | -89.6678
7.7333 | -89.6687
7.7336 | -89.6696
7.7339 | -89.6705
7.7343 | -89.6715
7.7346 | -89.6729
7.7349 | -89.6745
7.7352 | -89.6757
7.7355 | -89.6769
7.7358 | -89.6791
7.7361 | -89.6818
7.7364 | -89.6831
7.7367 | -89.684
7.737 | -89.6849
7.7374 | -89.6851
7.7377 | -89.6859
7.738 | -89.6864
7.7383 | -89.6878
7.7386 | -89.6888
7.7389 | -89.6893
7.7392 | -89.6896
7.7395 | -89.6896
7.7398 | -89.6902
7.7402 | -89.6911
7.7405 | -89.6924
7.7408 | -89.694
7.7411 | -89.6951
7.7414 | -89.6959
7.7417 | -89.6963
7.742 | -89.6972
7.7423 | -89.6979
7.7426 | -89.6989
7.7429 | -89.7005
7.7433 | -89.7017
7.7436 | -89.7026
7.7439 | -89.7038
7.7442 | -89.7048
7.7445 | -89.7056
7.7448 | -89.7066
7.7451 | -89.7079
7.7454 | -89.7093
7.7457 | -89.7106
7.7461 | -89.7113
7.7464 | -89.712
7.7467 | -89.713
7.747 | -89.7142
7.7473 | -89.715
7.7476 | -89.7159
7.7479 | -89.7173
7.7482 | -89.7188
7.7485 | -89.72
7.7488 | -89.7209
7.7492 | -89.7212
7.7495 | -89.7219
7.7498 | -89.7233
7.7501 | -89.7245
7.7504 | -89.7252
7.7507 | -89.726
7.751 | -89.7268
7.7513 | -89.7275
7.7516 | -89.7283
7.752 | -89.7292
7.7523 | -89.7301
7.7526 | -89.7301
7.7529 | -89.7305
7.7532 | -89.7315
7.7535 | -89.7327
7.7538 | -89.7337
7.7541 | -89.7347
7.7544 | -89.7355
7.7547 | -89.7365
7.7551 | -89.7379
7.7554 | -89.7388
7.7557 | -89.7393
7.756 | -89.74
7.7563 | -89.7407
7.7566 | -89.7415
7.7569 | -89.7426
7.7572 | -89.7434
7.7575 | -89.7446
7.7579 | -89.7459
7.7582 | -89.747
7.7585 | -89.7478
7.7588 | -89.749
7.7591 | -89.7506
7.7594 | -89.7528
7.7597 | -89.7546
7.76 | -89.7555
7.7603 | -89.7564
7.7607 | -89.7573
7.761 | -89.7595
7.7613 | -89.7607
7.7616 | -89.7618
7.7619 | -89.7632
7.7622 | -89.7644
7.7625 | -89.7656
7.7628 | -89.7673
7.7631 | -89.769
7.7634 | -89.7706
7.7638 | -89.7717
7.7641 | -89.7726
7.7644 | -89.7739
7.7647 | -89.7756
7.765 | -89.7767
7.7653 | -89.7778
7.7656 | -89.7784
7.7659 | -89.7794
7.7662 | -89.7806
7.7665 | -89.7818
7.7669 | -89.7828
7.7672 | -89.7837
7.7675 | -89.7849
7.7678 | -89.7864
7.7681 | -89.7876
7.7684 | -89.7888
7.7687 | -89.7896
7.769 | -89.791
7.7693 | -89.7925
7.7697 | -89.7934
7.77 | -89.7946
7.7703 | -89.7957
7.7706 | -89.7967
7.7709 | -89.7977
7.7712 | -89.7992
7.7715 | -89.8002
7.7718 | -89.8011
7.7721 | -89.8019
7.7725 | -89.8028
7.7728 | -89.804
7.7731 | -89.805
7.7734 | -89.8062
7.7737 | -89.8076
7.774 | -89.8087
7.7743 | -89.8101
7.7746 | -89.8113
7.7749 | -89.8126
7.7752 | -89.8141
7.7756 | -89.8157
7.7759 | -89.8169
7.7762 | -89.8179
7.7765 | -89.8194
7.7768 | -89.8207
7.7771 | -89.8227
7.7774 | -89.8247
7.7777 | -89.8263
7.778 | -89.8271
7.7784 | -89.8284
7.7787 | -89.8301
7.779 | -89.8315
7.7793 | -89.8332
7.7796 | -89.8339
7.7799 | -89.8348
7.7802 | -89.8358
7.7805 | -89.8373
7.7808 | -89.8394
7.7811 | -89.8413
7.7815 | -89.8426
7.7818 | -89.8431
7.7821 | -89.844
7.7824 | -89.8433
7.7827 | -89.8439
7.783 | -89.8447
7.7833 | -89.845
7.7836 | -89.846
7.7839 | -89.847
7.7843 | -89.8473
7.7846 | -89.8482
7.7849 | -89.8494
7.7852 | -89.8504
7.7855 | -89.8509
7.7858 | -89.8517
7.7861 | -89.8529
7.7864 | -89.8537
7.7867 | -89.8542
7.787 | -89.8549
7.7874 | -89.8559
7.7877 | -89.8566
7.788 | -89.857
7.7883 | -89.8583
7.7886 | -89.8595
7.7889 | -89.8604
7.7892 | -89.8615
7.7895 | -89.8626
7.7898 | -89.8636
7.7902 | -89.8647
7.7905 | -89.8656
7.7908 | -89.8664
7.7911 | -89.8677
7.7914 | -89.8693
7.7917 | -89.8704
7.792 | -89.8715
7.7923 | -89.8727
7.7926 | -89.8731
7.7929 | -89.8737
7.7933 | -89.8749
7.7936 | -89.8761
7.7939 | -89.8772
7.7942 | -89.878
7.7945 | -89.8801
7.7948 | -89.881
7.7951 | -89.8827
7.7954 | -89.8839
7.7957 | -89.885
7.7961 | -89.8861
7.7964 | -89.8875
7.7967 | -89.8891
7.797 | -89.89
7.7973 | -89.891
7.7976 | -89.8923
7.7979 | -89.8937
7.7982 | -89.8949
7.7985 | -89.8959
7.7988 | -89.8972
7.7992 | -89.8984
7.7995 | -89.8997
7.7998 | -89.9006
7.8001 | -89.9021
7.8004 | -89.9032
7.8007 | -89.9042
7.801 | -89.9064
7.8013 | -89.9074
7.8016 | -89.9073
7.802 | -89.9078
7.8023 | -89.9089
7.8026 | -89.9092
7.8029 | -89.91
7.8032 | -89.9111
7.8035 | -89.9125
7.8038 | -89.914
7.8041 | -89.9154
7.8044 | -89.9167
7.8047 | -89.9178
7.8051 | -89.9188
7.8054 | -89.9199
7.8057 | -89.9209
7.806 | -89.9221
7.8063 | -89.9231
7.8066 | -89.9243
7.8069 | -89.9257
7.8072 | -89.9273
7.8075 | -89.9281
7.8079 | -89.9285
7.8082 | -89.9293
7.8085 | -89.93
7.8088 | -89.9309
7.8091 | -89.9319
7.8094 | -89.9331
7.8097 | -89.9347
7.81 | -89.9362
7.8103 | -89.9378
7.8106 | -89.9388
7.811 | -89.9398
7.8113 | -89.9412
7.8116 | -89.9433
7.8119 | -89.945
7.8122 | -89.9461
7.8125 | -89.9477
7.8128 | -89.949
7.8131 | -89.9502
7.8134 | -89.9513
7.8138 | -89.9523
7.8141 | -89.9534
7.8144 | -89.9543
7.8147 | -89.9544
7.815 | -89.9553
7.8153 | -89.9569
7.8156 | -89.9583
7.8159 | -89.9596
7.8162 | -89.9609
7.8165 | -89.9622
7.8169 | -89.963
7.8172 | -89.9641
7.8175 | -89.9653
7.8178 | -89.9664
7.8181 | -89.9674
7.8184 | -89.9685
7.8187 | -89.9699
7.819 | -89.971
7.8193 | -89.9721
7.8197 | -89.973
7.82 | -89.9739
7.8203 | -89.9749
7.8206 | -89.9755
7.8209 | -89.9765
7.8212 | -89.9777
7.8215 | -89.9787
7.8218 | -89.9798
7.8221 | -89.9804
7.8224 | -89.9813
7.8228 | -89.9822
7.8231 | -89.9829
7.8234 | -89.9836
7.8237 | -89.985
7.824 | -89.9867
7.8243 | -89.9884
7.8246 | -89.9895
7.8249 | -89.9903
7.8252 | -89.9911
7.8256 | -89.9938
7.8259 | -89.9969
7.8262 | -89.9988
7.8265 | -90
7.8268 | -90.0012
7.8271 | -90.0027
7.8274 | -90.0043
7.8277 | -90.0049
7.828 | -90.0052
7.8283 | -90.0061
7.8287 | -90.0068
7.829 | -90.0077
7.8293 | -90.0079
7.8296 | -90.0093
7.8299 | -90.0103
7.8302 | -90.0117
7.8305 | -90.0123
7.8308 | -90.013
7.8311 | -90.0141
7.8315 | -90.0153
7.8318 | -90.0161
7.8321 | -90.0168
7.8324 | -90.0178
7.8327 | -90.0188
7.833 | -90.02
7.8333 | -90.0213
7.8336 | -90.0225
7.8339 | -90.0232
7.8342 | -90.0239
7.8346 | -90.0248
7.8349 | -90.0276
7.8352 | -90.0296
7.8355 | -90.0306
7.8358 | -90.0315
7.8361 | -90.0331
7.8364 | -90.0343
7.8367 | -90.0349
7.837 | -90.0354
7.8374 | -90.0363
7.8377 | -90.0371
7.838 | -90.0378
7.8383 | -90.0386
7.8386 | -90.0407
7.8389 | -90.0423
7.8392 | -90.0432
7.8395 | -90.0433
7.8398 | -90.0443
7.8401 | -90.0452
7.8405 | -90.0459
7.8408 | -90.0466
7.8411 | -90.0471
7.8414 | -90.0481
7.8417 | -90.049
7.842 | -90.0499
7.8423 | -90.0508
7.8426 | -90.0516
7.8429 | -90.0525
7.8433 | -90.0537
7.8436 | -90.0548
7.8439 | -90.0557
7.8442 | -90.0569
7.8445 | -90.0581
7.8448 | -90.059
7.8451 | -90.0609
7.8454 | -90.0622
7.8457 | -90.0628
7.846 | -90.0634
7.8464 | -90.0644
7.8467 | -90.0652
7.847 | -90.0663
7.8473 | -90.0672
7.8476 | -90.0681
7.8479 | -90.0689
7.8482 | -90.0697
7.8485 | -90.0705
7.8488 | -90.0716
7.8492 | -90.0726
7.8495 | -90.0734
7.8498 | -90.0744
7.8501 | -90.0749
7.8504 | -90.0759
7.8507 | -90.0765
7.851 | -90.0774
7.8513 | -90.0782
7.8516 | -90.079
7.8519 | -90.0803
7.8523 | -90.0817
7.8526 | -90.0826
7.8529 | -90.0833
7.8532 | -90.084
7.8535 | -90.0851
7.8538 | -90.0863
7.8541 | -90.0875
7.8544 | -90.088
7.8547 | -90.0893
7.8551 | -90.0906
7.8554 | -90.092
7.8557 | -90.0931
7.856 | -90.0939
7.8563 | -90.0948
7.8566 | -90.0959
7.8569 | -90.0965
7.8572 | -90.0972
7.8575 | -90.0984
7.8578 | -90.0995
7.8582 | -90.1005
7.8585 | -90.1017
7.8588 | -90.1029
7.8591 | -90.1036
7.8594 | -90.1048
7.8597 | -90.106
7.86 | -90.1068
7.8603 | -90.1078
7.8606 | -90.1077
7.861 | -90.1079
7.8613 | -90.1093
7.8616 | -90.1107
7.8619 | -90.1119
7.8622 | -90.1126
7.8625 | -90.1134
7.8628 | -90.1141
7.8631 | -90.1153
7.8634 | -90.1164
7.8637 | -90.1176
7.8641 | -90.1192
7.8644 | -90.1204
7.8647 | -90.1209
7.865 | -90.1218
7.8653 | -90.1235
7.8656 | -90.125
7.8659 | -90.1264
7.8662 | -90.1274
7.8665 | -90.1273
7.8669 | -90.1282
7.8672 | -90.1291
7.8675 | -90.1296
7.8678 | -90.1299
7.8681 | -90.1306
7.8684 | -90.1315
7.8687 | -90.1322
7.869 | -90.1332
7.8693 | -90.1344
7.8696 | -90.1353
7.87 | -90.136
7.8703 | -90.1368
7.8706 | -90.138
7.8709 | -90.139
7.8712 | -90.1396
7.8715 | -90.1402
7.8718 | -90.141
7.8721 | -90.1419
7.8724 | -90.143
7.8728 | -90.1438
7.8731 | -90.1451
7.8734 | -90.1462
7.8737 | -90.1469
7.874 | -90.148
7.8743 | -90.149
7.8746 | -90.1499
7.8749 | -90.1508
7.8752 | -90.1521
7.8755 | -90.1531
7.8759 | -90.1541
7.8762 | -90.1546
7.8765 | -90.1549
7.8768 | -90.1559
7.8771 | -90.1569
7.8774 | -90.1577
7.8777 | -90.1585
7.878 | -90.1594
7.8783 | -90.1604
7.8787 | -90.1614
7.879 | -90.1627
7.8793 | -90.1637
7.8796 | -90.1647
7.8799 | -90.1665
7.8802 | -90.1682
7.8805 | -90.1691
7.8808 | -90.1699
7.8811 | -90.1707
7.8814 | -90.1716
7.8818 | -90.1728
7.8821 | -90.1735
7.8824 | -90.1745
7.8827 | -90.1754
7.883 | -90.1758
7.8833 | -90.176
7.8836 | -90.1767
7.8839 | -90.1776
7.8842 | -90.1784
7.8846 | -90.1795
7.8849 | -90.1804
7.8852 | -90.1811
7.8855 | -90.1822
7.8858 | -90.1834
7.8861 | -90.1832
7.8864 | -90.1832
7.8867 | -90.1839
7.887 | -90.1852
7.8873 | -90.1874
7.8877 | -90.188
7.888 | -90.1895
7.8883 | -90.1914
7.8886 | -90.1921
7.8889 | -90.1931
7.8892 | -90.1949
7.8895 | -90.1962
7.8898 | -90.1973
7.8901 | -90.1985
7.8905 | -90.1994
7.8908 | -90.2004
7.8911 | -90.201
7.8914 | -90.202
7.8917 | -90.2022
7.892 | -90.2025
7.8923 | -90.2031
7.8926 | -90.2034
7.8929 | -90.2041
7.8933 | -90.2048
7.8936 | -90.2057
7.8939 | -90.2065
7.8942 | -90.2073
7.8945 | -90.2087
7.8948 | -90.2098
7.8951 | -90.2109
7.8954 | -90.2116
7.8957 | -90.2123
7.896 | -90.213
7.8964 | -90.2135
7.8967 | -90.2141
7.897 | -90.2146
7.8973 | -90.2157
7.8976 | -90.2167
7.8979 | -90.2177
7.8982 | -90.2186
7.8985 | -90.2188
7.8988 | -90.219
7.8991 | -90.2193
7.8995 | -90.2199
7.8998 | -90.2213
7.9001 | -90.2212
7.9004 | -90.2214
7.9007 | -90.2222
7.901 | -90.2237
7.9013 | -90.2246
7.9016 | -90.2251
7.9019 | -90.2251
7.9023 | -90.2254
7.9026 | -90.226
7.9029 | -90.2265
7.9032 | -90.2271
7.9035 | -90.2278
7.9038 | -90.2285
7.9041 | -90.2297
7.9044 | -90.2306
7.9047 | -90.2317
7.9051 | -90.2337
7.9054 | -90.2351
7.9057 | -90.2355
7.906 | -90.2357
7.9063 | -90.236
7.9066 | -90.236
7.9069 | -90.2366
7.9072 | -90.2376
7.9075 | -90.2388
7.9078 | -90.2394
7.9082 | -90.2392
7.9085 | -90.2395
7.9088 | -90.2403
7.9091 | -90.2408
7.9094 | -90.2421
7.9097 | -90.2424
7.91 | -90.243
7.9103 | -90.2436
7.9106 | -90.245
7.9109 | -90.2459
7.9113 | -90.2467
7.9116 | -90.2473
7.9119 | -90.2487
7.9122 | -90.2494
7.9125 | -90.2502
7.9128 | -90.251
7.9131 | -90.2517
7.9134 | -90.2521
7.9137 | -90.2526
7.9141 | -90.2532
7.9144 | -90.2542
7.9147 | -90.2553
7.915 | -90.2561
7.9153 | -90.2565
7.9156 | -90.2562
7.9159 | -90.2564
7.9162 | -90.2573
7.9165 | -90.2579
7.9169 | -90.2593
7.9172 | -90.26
7.9175 | -90.2612
7.9178 | -90.262
7.9181 | -90.2625
7.9184 | -90.2631
7.9187 | -90.2637
7.919 | -90.2646
7.9193 | -90.2653
7.9196 | -90.2659
7.92 | -90.2667
7.9203 | -90.2661
7.9206 | -90.2667
7.9209 | -90.2667
7.9212 | -90.2675
7.9215 | -90.2683
7.9218 | -90.2692
7.9221 | -90.2699
7.9224 | -90.2702
7.9228 | -90.2708
7.9231 | -90.2715
7.9234 | -90.2718
7.9237 | -90.2722
7.924 | -90.2732
7.9243 | -90.2742
7.9246 | -90.2744
7.9249 | -90.2751
7.9252 | -90.276
7.9255 | -90.2768
7.9259 | -90.2775
7.9262 | -90.2781
7.9265 | -90.2785
7.9268 | -90.2793
7.9271 | -90.2801
7.9274 | -90.2806
7.9277 | -90.2808
7.928 | -90.2816
7.9283 | -90.2829
7.9287 | -90.2835
7.929 | -90.2841
7.9293 | -90.284
7.9296 | -90.2851
7.9299 | -90.2861
7.9302 | -90.2866
7.9305 | -90.2878
7.9308 | -90.2891
7.9311 | -90.2899
7.9314 | -90.2904
7.9318 | -90.2911
7.9321 | -90.2914
7.9324 | -90.2918
7.9327 | -90.2921
7.933 | -90.2924
7.9333 | -90.293
7.9336 | -90.2936
7.9339 | -90.2941
7.9342 | -90.2949
7.9346 | -90.2957
7.9349 | -90.2966
7.9352 | -90.2972
7.9355 | -90.2975
7.9358 | -90.2979
7.9361 | -90.2984
7.9364 | -90.2993
7.9367 | -90.2997
7.937 | -90.3004
7.9373 | -90.3009
7.9377 | -90.3013
7.938 | -90.3018
7.9383 | -90.3025
7.9386 | -90.303
7.9389 | -90.3044
7.9392 | -90.3054
7.9395 | -90.3062
7.9398 | -90.3064
7.9401 | -90.3071
7.9405 | -90.3082
7.9408 | -90.3088
7.9411 | -90.3095
7.9414 | -90.3105
7.9417 | -90.3115
7.942 | -90.3123
7.9423 | -90.3128
7.9426 | -90.3131
7.9429 | -90.3136
7.9432 | -90.3142
7.9436 | -90.315
7.9439 | -90.3151
7.9442 | -90.3153
7.9445 | -90.316
7.9448 | -90.3167
7.9451 | -90.3171
7.9454 | -90.3178
7.9457 | -90.3187
7.946 | -90.3195
7.9464 | -90.3209
7.9467 | -90.3235
7.947 | -90.3247
7.9473 | -90.3256
7.9476 | -90.3255
7.9479 | -90.325
7.9482 | -90.3264
7.9485 | -90.3278
7.9488 | -90.3288
7.9491 | -90.3292
7.9495 | -90.3295
7.9498 | -90.3295
7.9501 | -90.33
7.9504 | -90.3308
7.9507 | -90.3321
7.951 | -90.3335
7.9513 | -90.3341
7.9516 | -90.335
7.9519 | -90.3369
7.9523 | -90.3393
7.9526 | -90.34
7.9529 | -90.3405
7.9532 | -90.3414
7.9535 | -90.3418
7.9538 | -90.3421
7.9541 | -90.3431
7.9544 | -90.3436
7.9547 | -90.3443
7.955 | -90.3454
7.9554 | -90.3461
7.9557 | -90.3467
7.956 | -90.3474
7.9563 | -90.348
7.9566 | -90.3488
7.9569 | -90.3493
7.9572 | -90.3507
7.9575 | -90.3515
7.9578 | -90.3517
7.9582 | -90.3534
7.9585 | -90.3545
7.9588 | -90.355
7.9591 | -90.3555
7.9594 | -90.3564
7.9597 | -90.357
7.96 | -90.3573
7.9603 | -90.3577
7.9606 | -90.3589
7.9609 | -90.3597
7.9613 | -90.3602
7.9616 | -90.3606
7.9619 | -90.361
7.9622 | -90.3612
7.9625 | -90.3612
7.9628 | -90.3615
7.9631 | -90.3621
7.9634 | -90.3625
7.9637 | -90.3631
7.9641 | -90.3637
7.9644 | -90.3641
7.9647 | -90.3648
7.965 | -90.3647
7.9653 | -90.3651
7.9656 | -90.3655
7.9659 | -90.3661
7.9662 | -90.3669
7.9665 | -90.3676
7.9668 | -90.3679
7.9672 | -90.3684
7.9675 | -90.3693
7.9678 | -90.3698
7.9681 | -90.37
7.9684 | -90.3704
7.9687 | -90.3711
7.969 | -90.3723
7.9693 | -90.3731
7.9696 | -90.3735
7.97 | -90.3745
7.9703 | -90.3752
7.9706 | -90.3758
7.9709 | -90.3765
7.9712 | -90.3777
7.9715 | -90.3788
7.9718 | -90.3787
7.9721 | -90.3781
7.9724 | -90.3793
7.9727 | -90.381
7.9731 | -90.3826
7.9734 | -90.3836
7.9737 | -90.3837
7.974 | -90.384
7.9743 | -90.3843
7.9746 | -90.385
7.9749 | -90.3858
7.9752 | -90.3862
7.9755 | -90.3868
7.9759 | -90.3878
7.9762 | -90.389
7.9765 | -90.3894
7.9768 | -90.389
7.9771 | -90.3895
7.9774 | -90.3898
7.9777 | -90.3902
7.978 | -90.3907
7.9783 | -90.3911
7.9786 | -90.3913
7.979 | -90.3918
7.9793 | -90.3924
7.9796 | -90.3928
7.9799 | -90.3932
7.9802 | -90.3935
7.9805 | -90.3942
7.9808 | -90.3947
7.9811 | -90.3952
7.9814 | -90.3956
7.9818 | -90.3961
7.9821 | -90.3966
7.9824 | -90.3969
7.9827 | -90.3973
7.983 | -90.398
7.9833 | -90.3985
7.9836 | -90.3999
7.9839 | -90.4012
7.9842 | -90.4019
7.9845 | -90.4025
7.9849 | -90.403
7.9852 | -90.4037
7.9855 | -90.4044
7.9858 | -90.4048
7.9861 | -90.4054
7.9864 | -90.4059
7.9867 | -90.4069
7.987 | -90.4083
7.9873 | -90.4089
7.9877 | -90.4091
7.988 | -90.4093
7.9883 | -90.4103
7.9886 | -90.411
7.9889 | -90.4112
7.9892 | -90.4111
7.9895 | -90.4115
7.9898 | -90.4115
7.9901 | -90.4112
7.9904 | -90.4116
7.9908 | -90.4121
7.9911 | -90.4121
7.9914 | -90.4119
7.9917 | -90.4122
7.992 | -90.4131
7.9923 | -90.4135
7.9926 | -90.4139
7.9929 | -90.4143
7.9932 | -90.4146
7.9936 | -90.4145
7.9939 | -90.4153
7.9942 | -90.4171
7.9945 | -90.4178
7.9948 | -90.4183
7.9951 | -90.4189
7.9954 | -90.4193
7.9957 | -90.4196
7.996 | -90.4201
7.9963 | -90.4205
7.9967 | -90.4215
7.997 | -90.4226
7.9973 | -90.4238
7.9976 | -90.4252
7.9979 | -90.426
7.9982 | -90.4269
7.9985 | -90.4273
7.9988 | -90.4276
7.9991 | -90.4281
7.9995 | -90.4286
7.9998 | -90.4289
8.0001 | -90.4294
8.0004 | -90.4298
8.0007 | -90.4307
8.001 | -90.4313
8.0013 | -90.4317
8.0016 | -90.4321
8.0019 | -90.4325
8.0022 | -90.4335
8.0026 | -90.4342
8.0029 | -90.4347
8.0032 | -90.4353
8.0035 | -90.4359
8.0038 | -90.4367
8.0041 | -90.4375
8.0044 | -90.4388
8.0047 | -90.4389
8.005 | -90.439
8.0054 | -90.4397
8.0057 | -90.4419
8.006 | -90.4441
8.0063 | -90.4449
8.0066 | -90.4451
8.0069 | -90.446
8.0072 | -90.4463
8.0075 | -90.4468
8.0078 | -90.4472
8.0081 | -90.4474
8.0085 | -90.4479
8.0088 | -90.4483
8.0091 | -90.4486
8.0094 | -90.4491
8.0097 | -90.4491
8.01 | -90.4497
8.0103 | -90.4499
8.0106 | -90.4504
8.0109 | -90.4517
8.0113 | -90.4526
8.0116 | -90.4534
8.0119 | -90.4537
8.0122 | -90.4543
8.0125 | -90.4548
8.0128 | -90.4554
8.0131 | -90.4564
8.0134 | -90.4573
8.0137 | -90.458
8.014 | -90.4585
8.0144 | -90.4588
8.0147 | -90.4593
8.015 | -90.4597
8.0153 | -90.4599
8.0156 | -90.4607
8.0159 | -90.4618
8.0162 | -90.4621
8.0165 | -90.4624
8.0168 | -90.4627
8.0172 | -90.464
8.0175 | -90.4666
8.0178 | -90.4676
8.0181 | -90.4682
8.0184 | -90.4684
8.0187 | -90.4686
8.019 | -90.4693
8.0193 | -90.4701
8.0196 | -90.4707
8.0199 | -90.4715
8.0203 | -90.472
8.0206 | -90.4724
8.0209 | -90.473
8.0212 | -90.4736
8.0215 | -90.4743
8.0218 | -90.4749
8.0221 | -90.4759
8.0224 | -90.4762
8.0227 | -90.4765
8.0231 | -90.4772
8.0234 | -90.4778
8.0237 | -90.478
8.024 | -90.4788
8.0243 | -90.4795
8.0246 | -90.4804
8.0249 | -90.4815
8.0252 | -90.4828
8.0255 | -90.4842
8.0258 | -90.4845
8.0262 | -90.4848
8.0265 | -90.4855
8.0268 | -90.4865
8.0271 | -90.4873
8.0274 | -90.4883
8.0277 | -90.4882
8.028 | -90.4875
8.0283 | -90.4873
8.0286 | -90.4874
8.029 | -90.4878
8.0293 | -90.4884
8.0296 | -90.4888
8.0299 | -90.4891
8.0302 | -90.489
8.0305 | -90.4889
8.0308 | -90.4894
8.0311 | -90.4899
8.0314 | -90.4904
8.0317 | -90.491
8.0321 | -90.491
8.0324 | -90.4911
8.0327 | -90.4914
8.033 | -90.492
8.0333 | -90.4929
8.0336 | -90.4928
8.0339 | -90.4924
8.0342 | -90.4929
8.0345 | -90.4945
8.0349 | -90.4952
8.0352 | -90.4945
8.0355 | -90.494
8.0358 | -90.4946
8.0361 | -90.4951
8.0364 | -90.4957
8.0367 | -90.4967
8.037 | -90.4974
8.0373 | -90.4981
8.0376 | -90.4982
8.038 | -90.4983
8.0383 | -90.4994
8.0386 | -90.5014
8.0389 | -90.5022
8.0392 | -90.5026
8.0395 | -90.5035
8.0398 | -90.5044
8.0401 | -90.5049
8.0404 | -90.5058
8.0408 | -90.5069
8.0411 | -90.5078
8.0414 | -90.5084
8.0417 | -90.5087
8.042 | -90.5092
8.0423 | -90.5097
8.0426 | -90.5103
8.0429 | -90.5111
8.0432 | -90.5116
8.0435 | -90.5121
8.0439 | -90.5124
8.0442 | -90.5127
8.0445 | -90.5132
8.0448 | -90.5146
8.0451 | -90.5163
8.0454 | -90.5174
8.0457 | -90.5186
8.046 | -90.5188
8.0463 | -90.5192
8.0467 | -90.5205
8.047 | -90.5208
8.0473 | -90.5213
8.0476 | -90.5222
8.0479 | -90.5225
8.0482 | -90.5227
8.0485 | -90.5232
8.0488 | -90.5235
8.0491 | -90.5238
8.0495 | -90.5249
8.0498 | -90.5263
8.0501 | -90.5269
8.0504 | -90.5279
8.0507 | -90.5286
8.051 | -90.5309
8.0513 | -90.5335
8.0516 | -90.5343
8.0519 | -90.535
8.0522 | -90.5356
8.0526 | -90.5362
8.0529 | -90.5365
8.0532 | -90.5368
8.0535 | -90.5386
8.0538 | -90.5413
8.0541 | -90.5425
8.0544 | -90.5419
8.0547 | -90.5417
8.055 | -90.5421
8.0554 | -90.5427
8.0557 | -90.5434
8.056 | -90.5438
8.0563 | -90.5443
8.0566 | -90.5447
8.0569 | -90.5449
8.0572 | -90.5453
8.0575 | -90.5458
8.0578 | -90.5469
8.0581 | -90.5482
8.0585 | -90.5493
8.0588 | -90.5498
8.0591 | -90.5502
8.0594 | -90.5508
8.0597 | -90.5523
8.06 | -90.553
8.0603 | -90.5534
8.0606 | -90.5545
8.0609 | -90.5555
8.0612 | -90.5561
8.0616 | -90.5567
8.0619 | -90.5576
8.0622 | -90.5581
8.0625 | -90.5589
8.0628 | -90.5594
8.0631 | -90.56
8.0634 | -90.5605
8.0637 | -90.5615
8.064 | -90.5622
8.0644 | -90.5629
8.0647 | -90.5634
8.065 | -90.564
8.0653 | -90.5653
8.0656 | -90.5661
8.0659 | -90.5665
8.0662 | -90.5672
8.0665 | -90.5677
8.0668 | -90.568
8.0672 | -90.5686
8.0675 | -90.5691
8.0678 | -90.5695
8.0681 | -90.5701
8.0684 | -90.5717
8.0687 | -90.5724
8.069 | -90.5729
8.0693 | -90.5738
8.0696 | -90.5746
8.0699 | -90.5752
8.0703 | -90.5749
8.0706 | -90.5751
8.0709 | -90.5758
8.0712 | -90.5766
8.0715 | -90.577
8.0718 | -90.5777
8.0721 | -90.5783
8.0724 | -90.5792
8.0727 | -90.5795
8.0731 | -90.5796
8.0734 | -90.5803
8.0737 | -90.5809
8.074 | -90.5813
8.0743 | -90.5822
8.0746 | -90.5832
8.0749 | -90.5834
8.0752 | -90.5835
8.0755 | -90.5841
8.0758 | -90.5848
8.0762 | -90.5853
8.0765 | -90.586
8.0768 | -90.5872
8.0771 | -90.5879
8.0774 | -90.5884
8.0777 | -90.589
8.078 | -90.5889
8.0783 | -90.5891
8.0786 | -90.5897
8.079 | -90.5898
8.0793 | -90.5907
8.0796 | -90.5915
8.0799 | -90.5921
8.0802 | -90.5929
8.0805 | -90.5932
8.0808 | -90.5935
8.0811 | -90.5937
8.0814 | -90.5946
8.0817 | -90.595
8.0821 | -90.5957
8.0824 | -90.5964
8.0827 | -90.5972
8.083 | -90.5981
8.0833 | -90.599
8.0836 | -90.5998
8.0839 | -90.5998
8.0842 | -90.6003
8.0845 | -90.6018
8.0849 | -90.6025
8.0852 | -90.6033
8.0855 | -90.6036
8.0858 | -90.6036
8.0861 | -90.6036
8.0864 | -90.6036
8.0867 | -90.6039
8.087 | -90.6044
8.0873 | -90.6049
8.0876 | -90.6059
8.088 | -90.6061
8.0883 | -90.6066
8.0886 | -90.6073
8.0889 | -90.6079
8.0892 | -90.6085
8.0895 | -90.609
8.0898 | -90.6096
8.0901 | -90.6105
8.0904 | -90.6116
8.0908 | -90.6119
8.0911 | -90.6125
8.0914 | -90.6133
8.0917 | -90.6141
8.092 | -90.6138
8.0923 | -90.6141
8.0926 | -90.6145
8.0929 | -90.6151
8.0932 | -90.6155
8.0935 | -90.6155
8.0939 | -90.6159
8.0942 | -90.6165
8.0945 | -90.6172
8.0948 | -90.618
8.0951 | -90.6187
8.0954 | -90.6193
8.0957 | -90.6201
8.096 | -90.6207
8.0963 | -90.622
8.0967 | -90.6226
8.097 | -90.6223
8.0973 | -90.623
8.0976 | -90.6236
8.0979 | -90.6247
8.0982 | -90.6257
8.0985 | -90.6265
8.0988 | -90.6275
8.0991 | -90.6288
8.0994 | -90.63
8.0998 | -90.6309
8.1001 | -90.6316
8.1004 | -90.6321
8.1007 | -90.6325
8.101 | -90.6329
8.1013 | -90.6335
8.1016 | -90.6342
8.1019 | -90.6345
8.1022 | -90.6352
8.1026 | -90.6356
8.1029 | -90.6354
8.1032 | -90.6354
8.1035 | -90.6357
8.1038 | -90.6362
8.1041 | -90.637
8.1044 | -90.6373
8.1047 | -90.6378
8.105 | -90.6388
8.1053 | -90.6398
8.1057 | -90.6412
8.106 | -90.6426
8.1063 | -90.6432
8.1066 | -90.6439
8.1069 | -90.6442
8.1072 | -90.6445
8.1075 | -90.6449
8.1078 | -90.6458
8.1081 | -90.647
8.1085 | -90.6476
8.1088 | -90.6486
8.1091 | -90.6492
8.1094 | -90.6498
8.1097 | -90.6495
8.11 | -90.6499
8.1103 | -90.6506
8.1106 | -90.6512
8.1109 | -90.652
8.1112 | -90.653
8.1116 | -90.6537
8.1119 | -90.6543
8.1122 | -90.6549
8.1125 | -90.6553
8.1128 | -90.6562
8.1131 | -90.6568
8.1134 | -90.6575
8.1137 | -90.6582
8.114 | -90.6586
8.1144 | -90.6595
8.1147 | -90.6604
8.115 | -90.6611
8.1153 | -90.6616
8.1156 | -90.6619
8.1159 | -90.6629
8.1162 | -90.6635
8.1165 | -90.6642
8.1168 | -90.665
8.1171 | -90.6662
8.1175 | -90.6674
8.1178 | -90.6685
8.1181 | -90.6697
8.1184 | -90.6713
8.1187 | -90.6718
8.119 | -90.6728
8.1193 | -90.6734
8.1196 | -90.6748
8.1199 | -90.6755
8.1203 | -90.6758
8.1206 | -90.6766
8.1209 | -90.6773
8.1212 | -90.6781
8.1215 | -90.6792
8.1218 | -90.6807
8.1221 | -90.6814
8.1224 | -90.682
8.1227 | -90.6827
8.123 | -90.6836
8.1234 | -90.6849
8.1237 | -90.6856
8.124 | -90.6862
8.1243 | -90.6868
8.1246 | -90.6876
8.1249 | -90.6884
8.1252 | -90.6891
8.1255 | -90.6898
8.1258 | -90.6905
8.1262 | -90.6912
8.1265 | -90.6922
8.1268 | -90.6929
8.1271 | -90.6939
8.1274 | -90.6954
8.1277 | -90.6967
8.128 | -90.6976
8.1283 | -90.6984
8.1286 | -90.6989
8.1289 | -90.6988
8.1293 | -90.6996
8.1296 | -90.6997
8.1299 | -90.7001
8.1302 | -90.7011
8.1305 | -90.7017
8.1308 | -90.7022
8.1311 | -90.7034
8.1314 | -90.7043
8.1317 | -90.7051
8.1321 | -90.7047
8.1324 | -90.7053
8.1327 | -90.7059
8.133 | -90.7059
8.1333 | -90.7066
8.1336 | -90.707
8.1339 | -90.7071
8.1342 | -90.707
8.1345 | -90.7076
8.1348 | -90.7076
8.1352 | -90.7059
8.1355 | -90.7055
8.1358 | -90.7048
8.1361 | -90.7049
8.1364 | -90.7054
8.1367 | -90.7055
8.137 | -90.705
8.1373 | -90.7056
8.1376 | -90.7055
8.138 | -90.7036
8.1383 | -90.704
8.1386 | -90.7051
8.1389 | -90.7054
8.1392 | -90.707
8.1395 | -90.7076
8.1398 | -90.7085
8.1401 | -90.7089
8.1404 | -90.7095
8.1407 | -90.7095
8.1411 | -90.7093
8.1414 | -90.7094
8.1417 | -90.7102
8.142 | -90.7114
8.1423 | -90.7127
8.1426 | -90.7149
8.1429 | -90.7157
8.1432 | -90.7159
8.1435 | -90.7166
8.1439 | -90.7172
8.1442 | -90.7168
8.1445 | -90.7173
8.1448 | -90.7182
8.1451 | -90.7193
8.1454 | -90.7203
8.1457 | -90.7209
8.146 | -90.7213
8.1463 | -90.7222
8.1466 | -90.7225
8.147 | -90.7231
8.1473 | -90.7232
8.1476 | -90.7239
8.1479 | -90.725
8.1482 | -90.7262
8.1485 | -90.727
8.1488 | -90.7277
8.1491 | -90.7286
8.1494 | -90.7296
8.1498 | -90.7289
8.1501 | -90.7286
8.1504 | -90.7295
8.1507 | -90.73
8.151 | -90.7301
8.1513 | -90.7299
8.1516 | -90.7304
8.1519 | -90.7307
8.1522 | -90.7311
8.1525 | -90.7318
8.1529 | -90.7322
8.1532 | -90.7328
8.1535 | -90.7323
8.1538 | -90.732
8.1541 | -90.7327
8.1544 | -90.7333
8.1547 | -90.7339
8.155 | -90.7347
8.1553 | -90.7353
8.1557 | -90.7359
8.156 | -90.737
8.1563 | -90.7377
8.1566 | -90.7379
8.1569 | -90.7383
8.1572 | -90.739
8.1575 | -90.7402
8.1578 | -90.7416
8.1581 | -90.7424
8.1584 | -90.7429
8.1588 | -90.7436
8.1591 | -90.7445
8.1594 | -90.7459
8.1597 | -90.7469
8.16 | -90.7477
8.1603 | -90.7485
8.1606 | -90.7491
8.1609 | -90.7501
8.1612 | -90.7508
8.1616 | -90.7515
8.1619 | -90.7525
8.1622 | -90.7525
8.1625 | -90.754
8.1628 | -90.7549
8.1631 | -90.755
8.1634 | -90.7555
8.1637 | -90.756
8.164 | -90.7562
8.1643 | -90.7566
8.1647 | -90.7572
8.165 | -90.7579
8.1653 | -90.7585
8.1656 | -90.759
8.1659 | -90.7593
8.1662 | -90.7602
8.1665 | -90.7607
8.1668 | -90.7612
8.1671 | -90.7616
8.1675 | -90.7621
8.1678 | -90.7627
8.1681 | -90.7633
8.1684 | -90.7638
8.1687 | -90.7642
8.169 | -90.7644
8.1693 | -90.7653
8.1696 | -90.7657
8.1699 | -90.7663
8.1702 | -90.7673
8.1706 | -90.7686
8.1709 | -90.7698
8.1712 | -90.7705
8.1715 | -90.771
8.1718 | -90.7723
8.1721 | -90.7741
8.1724 | -90.775
8.1727 | -90.7755
8.173 | -90.776
8.1734 | -90.7764
8.1737 | -90.7769
8.174 | -90.7786
8.1743 | -90.7802
8.1746 | -90.7817
8.1749 | -90.7823
8.1752 | -90.7828
8.1755 | -90.7834
8.1758 | -90.784
8.1761 | -90.7839
8.1765 | -90.7837
8.1768 | -90.7841
8.1771 | -90.7848
8.1774 | -90.7858
8.1777 | -90.7866
8.178 | -90.787
8.1783 | -90.7873
8.1786 | -90.7882
8.1789 | -90.7889
8.1793 | -90.7894
8.1796 | -90.7898
8.1799 | -90.7912
8.1802 | -90.7918
8.1805 | -90.7924
8.1808 | -90.7922
8.1811 | -90.7927
8.1814 | -90.793
8.1817 | -90.7941
8.1821 | -90.7964
8.1824 | -90.7978
8.1827 | -90.7978
8.183 | -90.7976
8.1833 | -90.7984
8.1836 | -90.7983
8.1839 | -90.7981
8.1842 | -90.7986
8.1845 | -90.7991
8.1848 | -90.7993
8.1852 | -90.7997
8.1855 | -90.7994
8.1858 | -90.7994
8.1861 | -90.7996
8.1864 | -90.8
8.1867 | -90.8002
8.187 | -90.8004
8.1873 | -90.8007
8.1876 | -90.8021
8.188 | -90.8037
8.1883 | -90.8051
8.1886 | -90.8059
8.1889 | -90.8064
8.1892 | -90.8069
8.1895 | -90.8073
8.1898 | -90.8078
8.1901 | -90.8086
8.1904 | -90.8093
8.1907 | -90.8097
8.1911 | -90.8101
8.1914 | -90.8106
8.1917 | -90.811
8.192 | -90.8112
8.1923 | -90.8109
8.1926 | -90.8112
8.1929 | -90.8114
8.1932 | -90.8116
8.1935 | -90.8119
8.1938 | -90.8125
8.1942 | -90.8128
8.1945 | -90.8141
8.1948 | -90.8151
8.1951 | -90.8151
8.1954 | -90.816
8.1957 | -90.8166
8.196 | -90.8168
8.1963 | -90.8167
8.1966 | -90.816
8.197 | -90.8158
8.1973 | -90.8162
8.1976 | -90.817
8.1979 | -90.8176
8.1982 | -90.8176
8.1985 | -90.8181
8.1988 | -90.8184
8.1991 | -90.8189
8.1994 | -90.8195
8.1997 | -90.8198
8.2001 | -90.8201
8.2004 | -90.8205
8.2007 | -90.8211
8.201 | -90.8214
8.2013 | -90.8217
8.2016 | -90.8223
8.2019 | -90.8229
8.2022 | -90.8232
8.2025 | -90.8237
8.2029 | -90.824
8.2032 | -90.8245
8.2035 | -90.8247
8.2038 | -90.825
8.2041 | -90.8254
8.2044 | -90.8261
8.2047 | -90.8261
8.205 | -90.8263
8.2053 | -90.8267
8.2057 | -90.827
8.206 | -90.8272
8.2063 | -90.8271
8.2066 | -90.8278
8.2069 | -90.827
8.2072 | -90.8283
8.2075 | -90.829
8.2078 | -90.8291
8.2081 | -90.8288
8.2084 | -90.8284
8.2088 | -90.8282
8.2091 | -90.8285
8.2094 | -90.8285
8.2097 | -90.8276
8.21 | -90.8273
8.2103 | -90.8275
8.2106 | -90.8281
8.2109 | -90.8283
8.2112 | -90.8284
8.2116 | -90.8287
8.2119 | -90.8285
8.2122 | -90.8285
8.2125 | -90.8291
8.2128 | -90.8295
8.2131 | -90.8295
8.2134 | -90.8293
8.2137 | -90.8293
8.214 | -90.8291
8.2143 | -90.8291
8.2147 | -90.8296
8.215 | -90.8298
8.2153 | -90.8298
8.2156 | -90.8297
8.2159 | -90.8301
8.2162 | -90.8302
8.2165 | -90.8303
8.2168 | -90.8306
8.2171 | -90.8307
8.2175 | -90.8311
8.2178 | -90.8306
8.2181 | -90.831
8.2184 | -90.8313
8.2187 | -90.832
8.219 | -90.8329
8.2193 | -90.8332
8.2196 | -90.8331
8.2199 | -90.8333
8.2202 | -90.8336
8.2206 | -90.8338
8.2209 | -90.8339
8.2212 | -90.8338
8.2215 | -90.8341
8.2218 | -90.8356
8.2221 | -90.8368
8.2224 | -90.8375
8.2227 | -90.8379
8.223 | -90.8383
8.2234 | -90.8386
8.2237 | -90.8388
8.224 | -90.8389
8.2243 | -90.8385
8.2246 | -90.8386
8.2249 | -90.8389
8.2252 | -90.839
8.2255 | -90.8384
8.2258 | -90.8378
8.2261 | -90.838
8.2265 | -90.8383
8.2268 | -90.8377
8.2271 | -90.8374
8.2274 | -90.8373
8.2277 | -90.8383
8.228 | -90.8386
8.2283 | -90.8373
8.2286 | -90.8369
8.2289 | -90.8369
8.2293 | -90.8368
8.2296 | -90.8361
8.2299 | -90.8364
8.2302 | -90.8374
8.2305 | -90.838
8.2308 | -90.8383
8.2311 | -90.8386
8.2314 | -90.8387
8.2317 | -90.8385
8.232 | -90.8387
8.2324 | -90.8386
8.2327 | -90.8387
8.233 | -90.8387
8.2333 | -90.8387
8.2336 | -90.8388
8.2339 | -90.8381
8.2342 | -90.838
8.2345 | -90.8386
8.2348 | -90.8388
8.2352 | -90.839
8.2355 | -90.8391
8.2358 | -90.8394
8.2361 | -90.8395
8.2364 | -90.8397
8.2367 | -90.8398
8.237 | -90.8402
8.2373 | -90.8402
8.2376 | -90.8398
8.2379 | -90.8402
8.2383 | -90.8404
8.2386 | -90.8405
8.2389 | -90.8404
8.2392 | -90.8408
8.2395 | -90.8413
8.2398 | -90.8413
8.2401 | -90.8412
8.2404 | -90.8416
8.2407 | -90.842
8.2411 | -90.8422
8.2414 | -90.8424
8.2417 | -90.8426
8.242 | -90.8427
8.2423 | -90.8428
8.2426 | -90.843
8.2429 | -90.8427
8.2432 | -90.8427
8.2435 | -90.8423
8.2438 | -90.8423
8.2442 | -90.8425
8.2445 | -90.8426
8.2448 | -90.8429
8.2451 | -90.8429
8.2454 | -90.8429
8.2457 | -90.8429
8.246 | -90.8436
8.2463 | -90.8449
8.2466 | -90.8448
8.247 | -90.8439
8.2473 | -90.8433
8.2476 | -90.8435
8.2479 | -90.8441
8.2482 | -90.8442
8.2485 | -90.8443
8.2488 | -90.8452
8.2491 | -90.8465
8.2494 | -90.8466
8.2497 | -90.8471
8.2501 | -90.8478
8.2504 | -90.848
8.2507 | -90.8479
8.251 | -90.8477
8.2513 | -90.8477
8.2516 | -90.8476
8.2519 | -90.8465
8.2522 | -90.8463
8.2525 | -90.8459
8.2529 | -90.8452
8.2532 | -90.8446
8.2535 | -90.8444
8.2538 | -90.8442
8.2541 | -90.8445
8.2544 | -90.8445
8.2547 | -90.8445
8.255 | -90.8445
8.2553 | -90.8443
8.2556 | -90.8443
8.256 | -90.8445
8.2563 | -90.8445
8.2566 | -90.8444
8.2569 | -90.8447
8.2572 | -90.8451
8.2575 | -90.8448
8.2578 | -90.8443
8.2581 | -90.8443
8.2584 | -90.8442
8.2588 | -90.8441
8.2591 | -90.8441
8.2594 | -90.8449
8.2597 | -90.8454
8.26 | -90.8456
8.2603 | -90.8455
8.2606 | -90.8452
8.2609 | -90.8454
8.2612 | -90.8453
8.2615 | -90.8446
8.2619 | -90.8441
8.2622 | -90.8441
8.2625 | -90.8439
8.2628 | -90.8437
8.2631 | -90.8432
8.2634 | -90.8428
8.2637 | -90.8426
8.264 | -90.8429
8.2643 | -90.8426
8.2647 | -90.8419
8.265 | -90.8413
8.2653 | -90.8411
8.2656 | -90.8411
8.2659 | -90.8413
8.2662 | -90.841
8.2665 | -90.8402
8.2668 | -90.8397
8.2671 | -90.8389
8.2674 | -90.8378
8.2678 | -90.8373
8.2681 | -90.8373
8.2684 | -90.8374
8.2687 | -90.8376
8.269 | -90.8374
8.2693 | -90.8373
8.2696 | -90.8372
8.2699 | -90.837
8.2702 | -90.8375
8.2706 | -90.8377
8.2709 | -90.8374
8.2712 | -90.8374
8.2715 | -90.8373
8.2718 | -90.8372
8.2721 | -90.837
8.2724 | -90.8363
8.2727 | -90.8359
8.273 | -90.8357
8.2733 | -90.8356
8.2737 | -90.8355
8.274 | -90.8353
8.2743 | -90.8346
8.2746 | -90.8342
8.2749 | -90.8343
8.2752 | -90.8341
8.2755 | -90.8341
8.2758 | -90.834
8.2761 | -90.8332
8.2765 | -90.8325
8.2768 | -90.8324
8.2771 | -90.8321
8.2774 | -90.8321
8.2777 | -90.8319
8.278 | -90.8318
8.2783 | -90.8316
8.2786 | -90.8314
8.2789 | -90.8308
8.2792 | -90.8306
8.2796 | -90.8304
8.2799 | -90.8301
8.2802 | -90.8292
8.2805 | -90.8291
8.2808 | -90.8289
8.2811 | -90.8279
8.2814 | -90.8276
8.2817 | -90.8275
8.282 | -90.8275
8.2824 | -90.8279
8.2827 | -90.8283
8.283 | -90.8277
8.2833 | -90.8276
8.2836 | -90.8277
8.2839 | -90.8281
8.2842 | -90.8284
8.2845 | -90.8284
8.2848 | -90.8283
8.2851 | -90.8284
8.2855 | -90.8278
8.2858 | -90.8274
8.2861 | -90.8271
8.2864 | -90.8263
8.2867 | -90.8264
8.287 | -90.827
8.2873 | -90.8267
8.2876 | -90.8257
8.2879 | -90.8257
8.2883 | -90.8259
8.2886 | -90.8255
8.2889 | -90.8251
8.2892 | -90.825
8.2895 | -90.8242
8.2898 | -90.8234
8.2901 | -90.8228
8.2904 | -90.8222
8.2907 | -90.8218
8.291 | -90.8206
8.2914 | -90.8197
8.2917 | -90.8197
8.292 | -90.8202
8.2923 | -90.8204
8.2926 | -90.8203
8.2929 | -90.8195
8.2932 | -90.8177
8.2935 | -90.817
8.2938 | -90.8166
8.2942 | -90.8165
8.2945 | -90.8164
8.2948 | -90.8158
8.2951 | -90.8159
8.2954 | -90.8168
8.2957 | -90.8179
8.296 | -90.8174
8.2963 | -90.8167
8.2966 | -90.8166
8.2969 | -90.8165
8.2973 | -90.8161
8.2976 | -90.8159
8.2979 | -90.8149
8.2982 | -90.8147
8.2985 | -90.815
8.2988 | -90.8148
8.2991 | -90.8145
8.2994 | -90.8146
8.2997 | -90.814
8.3001 | -90.8137
8.3004 | -90.8135
8.3007 | -90.8132
8.301 | -90.8131
8.3013 | -90.8131
8.3016 | -90.8127
8.3019 | -90.8123
8.3022 | -90.8118
8.3025 | -90.8113
8.3028 | -90.8112
8.3032 | -90.8105
8.3035 | -90.81
8.3038 | -90.8099
8.3041 | -90.8098
8.3044 | -90.8097
8.3047 | -90.8094
8.305 | -90.8086
8.3053 | -90.808
8.3056 | -90.8066
8.306 | -90.8053
8.3063 | -90.8054
8.3066 | -90.8055
8.3069 | -90.8052
8.3072 | -90.8047
8.3075 | -90.8046
8.3078 | -90.8048
8.3081 | -90.8044
8.3084 | -90.8043
8.3087 | -90.8042
8.3091 | -90.8041
8.3094 | -90.8041
8.3097 | -90.8039
8.31 | -90.8044
8.3103 | -90.8044
8.3106 | -90.8032
8.3109 | -90.8011
8.3112 | -90.8001
8.3115 | -90.7998
8.3119 | -90.7993
8.3122 | -90.7989
8.3125 | -90.7983
8.3128 | -90.7975
8.3131 | -90.7978
8.3134 | -90.7979
8.3137 | -90.7968
8.314 | -90.7948
8.3143 | -90.7946
8.3146 | -90.7937
8.315 | -90.7929
8.3153 | -90.7923
8.3156 | -90.7917
8.3159 | -90.7911
8.3162 | -90.7908
8.3165 | -90.7907
8.3168 | -90.7902
8.3171 | -90.7897
8.3174 | -90.7894
8.3178 | -90.789
8.3181 | -90.7888
8.3184 | -90.7883
8.3187 | -90.7881
8.319 | -90.7871
8.3193 | -90.7859
8.3196 | -90.7849
8.3199 | -90.7843
8.3202 | -90.7837
8.3206 | -90.7831
8.3209 | -90.7827
8.3212 | -90.7822
8.3215 | -90.7814
8.3218 | -90.7808
8.3221 | -90.7805
8.3224 | -90.7798
8.3227 | -90.7797
8.323 | -90.7791
8.3233 | -90.7787
8.3237 | -90.7793
8.324 | -90.7795
8.3243 | -90.7793
8.3246 | -90.7792
8.3249 | -90.7788
8.3252 | -90.7783
8.3255 | -90.7779
8.3258 | -90.7774
8.3261 | -90.7766
8.3264 | -90.7762
8.3268 | -90.7761
8.3271 | -90.7753
8.3274 | -90.7753
8.3277 | -90.7749
8.328 | -90.7739
8.3283 | -90.7737
8.3286 | -90.7732
8.3289 | -90.7721
8.3292 | -90.7716
8.3296 | -90.7711
8.3299 | -90.7706
8.3302 | -90.7704
8.3305 | -90.7704
8.3308 | -90.7708
8.3311 | -90.7709
8.3314 | -90.7698
8.3317 | -90.7684
8.332 | -90.7679
8.3323 | -90.7678
8.3327 | -90.7672
8.333 | -90.7666
8.3333 | -90.7658
8.3336 | -90.7654
8.3339 | -90.765
8.3342 | -90.7644
8.3345 | -90.7642
8.3348 | -90.7638
8.3351 | -90.763
8.3355 | -90.7623
8.3358 | -90.7619
8.3361 | -90.7616
8.3364 | -90.7612
8.3367 | -90.7602
8.337 | -90.7593
8.3373 | -90.7591
8.3376 | -90.7584
8.3379 | -90.7569
8.3383 | -90.7561
8.3386 | -90.7556
8.3389 | -90.7553
8.3392 | -90.755
8.3395 | -90.7545
8.3398 | -90.7542
8.3401 | -90.7545
8.3404 | -90.755
8.3407 | -90.7544
8.341 | -90.7548
8.3414 | -90.7545
8.3417 | -90.7542
8.342 | -90.7545
8.3423 | -90.7533
8.3426 | -90.7531
8.3429 | -90.7532
8.3432 | -90.7528
8.3435 | -90.7525
8.3438 | -90.7528
8.3442 | -90.7523
8.3445 | -90.7516
8.3448 | -90.7509
8.3451 | -90.7506
8.3454 | -90.7506
8.3457 | -90.7502
8.346 | -90.7496
8.3463 | -90.749
8.3466 | -90.749
8.3469 | -90.7486
8.3473 | -90.7479
8.3476 | -90.748
8.3479 | -90.7479
8.3482 | -90.7469
8.3485 | -90.746
8.3488 | -90.7452
8.3491 | -90.7443
8.3494 | -90.742
8.3497 | -90.7405
8.3501 | -90.7396
8.3504 | -90.7384
8.3507 | -90.7378
8.351 | -90.7375
8.3513 | -90.7364
8.3516 | -90.7353
8.3519 | -90.7347
8.3522 | -90.7341
8.3525 | -90.7338
8.3528 | -90.7344
8.3532 | -90.7343
8.3535 | -90.7335
8.3538 | -90.7325
8.3541 | -90.7319
8.3544 | -90.7315
8.3547 | -90.7309
8.355 | -90.7307
8.3553 | -90.7304
8.3556 | -90.7303
8.356 | -90.73
8.3563 | -90.7295
8.3566 | -90.7288
8.3569 | -90.7278
8.3572 | -90.727
8.3575 | -90.7267
8.3578 | -90.7263
8.3581 | -90.7257
8.3584 | -90.7248
8.3587 | -90.7239
8.3591 | -90.7233
8.3594 | -90.7225
8.3597 | -90.7219
8.36 | -90.7211
8.3603 | -90.7205
8.3606 | -90.7199
8.3609 | -90.7182
8.3612 | -90.7176
8.3615 | -90.7177
8.3619 | -90.7171
8.3622 | -90.7171
8.3625 | -90.7177
8.3628 | -90.718
8.3631 | -90.7175
8.3634 | -90.7178
8.3637 | -90.7177
8.364 | -90.7172
8.3643 | -90.7161
8.3646 | -90.7153
8.365 | -90.7143
8.3653 | -90.7131
8.3656 | -90.7109
8.3659 | -90.7095
8.3662 | -90.7089
8.3665 | -90.7083
8.3668 | -90.7071
8.3671 | -90.706
8.3674 | -90.7053
8.3678 | -90.7045
8.3681 | -90.7032
8.3684 | -90.7023
8.3687 | -90.7016
8.369 | -90.7015
8.3693 | -90.7014
8.3696 | -90.7007
8.3699 | -90.7003
8.3702 | -90.6998
8.3705 | -90.6988
8.3709 | -90.6975
8.3712 | -90.6968
8.3715 | -90.696
8.3718 | -90.6948
8.3721 | -90.6942
8.3724 | -90.6939
8.3727 | -90.6933
8.373 | -90.6928
8.3733 | -90.6924
8.3737 | -90.6921
8.374 | -90.6918
8.3743 | -90.6913
8.3746 | -90.6908
8.3749 | -90.6906
8.3752 | -90.6899
8.3755 | -90.6894
8.3758 | -90.6891
8.3761 | -90.6881
8.3764 | -90.688
8.3768 | -90.6875
8.3771 | -90.6867
8.3774 | -90.6864
8.3777 | -90.6858
8.378 | -90.6853
8.3783 | -90.6847
8.3786 | -90.6844
8.3789 | -90.6843
8.3792 | -90.6839
8.3796 | -90.6832
8.3799 | -90.6826
8.3802 | -90.6816
8.3805 | -90.6807
8.3808 | -90.6793
8.3811 | -90.6782
8.3814 | -90.6776
8.3817 | -90.6772
8.382 | -90.6768
8.3823 | -90.6765
8.3827 | -90.6763
8.383 | -90.6759
8.3833 | -90.6754
8.3836 | -90.6748
8.3839 | -90.674
8.3842 | -90.6736
8.3845 | -90.673
8.3848 | -90.6726
8.3851 | -90.6723
8.3855 | -90.6726
8.3858 | -90.673
8.3861 | -90.6728
8.3864 | -90.6723
8.3867 | -90.6717
8.387 | -90.6708
8.3873 | -90.6702
8.3876 | -90.6696
8.3879 | -90.6692
8.3882 | -90.6685
8.3886 | -90.668
8.3889 | -90.6675
8.3892 | -90.6669
8.3895 | -90.6668
8.3898 | -90.6663
8.3901 | -90.6656
8.3904 | -90.6654
8.3907 | -90.6656
8.391 | -90.6651
8.3914 | -90.6647
8.3917 | -90.6642
8.392 | -90.6645
8.3923 | -90.664
8.3926 | -90.6635
8.3929 | -90.6633
8.3932 | -90.6634
8.3935 | -90.6635
8.3938 | -90.6631
8.3941 | -90.6624
8.3945 | -90.6619
8.3948 | -90.6611
8.3951 | -90.6605
8.3954 | -90.6603
8.3957 | -90.6598
8.396 | -90.6592
8.3963 | -90.6587
8.3966 | -90.6583
8.3969 | -90.6579
8.3973 | -90.6578
8.3976 | -90.658
8.3979 | -90.6585
8.3982 | -90.6588
8.3985 | -90.658
8.3988 | -90.6572
8.3991 | -90.6565
8.3994 | -90.6557
8.3997 | -90.6549
8.4 | -90.6548
8.4004 | -90.6547
8.4007 | -90.6544
8.401 | -90.6541
8.4013 | -90.6538
8.4016 | -90.6534
8.4019 | -90.6528
8.4022 | -90.6523
8.4025 | -90.6518
8.4028 | -90.6515
8.4032 | -90.6519
8.4035 | -90.6517
8.4038 | -90.6512
8.4041 | -90.6509
8.4044 | -90.6506
8.4047 | -90.6506
8.405 | -90.6504
8.4053 | -90.6501
8.4056 | -90.6499
8.4059 | -90.6497
8.4063 | -90.6491
8.4066 | -90.6486
8.4069 | -90.6483
8.4072 | -90.6472
8.4075 | -90.6462
8.4078 | -90.6451
8.4081 | -90.6441
8.4084 | -90.6433
8.4087 | -90.6427
8.4091 | -90.6421
8.4094 | -90.6414
8.4097 | -90.641
8.41 | -90.6408
8.4103 | -90.6405
8.4106 | -90.6401
8.4109 | -90.6395
8.4112 | -90.6391
8.4115 | -90.639
8.4118 | -90.6391
8.4122 | -90.6388
8.4125 | -90.6383
8.4128 | -90.6375
8.4131 | -90.6373
8.4134 | -90.6367
8.4137 | -90.6358
8.414 | -90.6348
8.4143 | -90.6347
8.4146 | -90.6346
8.415 | -90.6342
8.4153 | -90.634
8.4156 | -90.6336
8.4159 | -90.6334
8.4162 | -90.6328
8.4165 | -90.632
8.4168 | -90.632
8.4171 | -90.6313
8.4174 | -90.6304
8.4177 | -90.63
8.4181 | -90.6295
8.4184 | -90.629
8.4187 | -90.6283
8.419 | -90.6279
8.4193 | -90.6275
8.4196 | -90.6268
8.4199 | -90.6261
8.4202 | -90.6256
8.4205 | -90.6242
8.4209 | -90.6234
8.4212 | -90.6227
8.4215 | -90.6227
8.4218 | -90.6227
8.4221 | -90.6229
8.4224 | -90.6225
8.4227 | -90.6218
8.423 | -90.6212
8.4233 | -90.6206
8.4236 | -90.6201
8.424 | -90.6195
8.4243 | -90.6193
8.4246 | -90.6188
8.4249 | -90.6183
8.4252 | -90.6181
8.4255 | -90.6176
8.4258 | -90.6182
8.4261 | -90.6185
8.4264 | -90.6181
8.4268 | -90.6177
8.4271 | -90.6173
8.4274 | -90.6169
8.4277 | -90.6169
8.428 | -90.6165
8.4283 | -90.6161
8.4286 | -90.6158
8.4289 | -90.6158
8.4292 | -90.6158
8.4295 | -90.6154
8.4299 | -90.615
8.4302 | -90.6145
8.4305 | -90.614
8.4308 | -90.6138
8.4311 | -90.6137
8.4314 | -90.6134
8.4317 | -90.6132
8.432 | -90.6138
8.4323 | -90.6135
8.4327 | -90.6132
8.433 | -90.6128
8.4333 | -90.6127
8.4336 | -90.6129
8.4339 | -90.6124
8.4342 | -90.612
8.4345 | -90.6118
8.4348 | -90.6111
8.4351 | -90.6107
8.4354 | -90.6105
8.4358 | -90.6103
8.4361 | -90.6099
8.4364 | -90.6096
8.4367 | -90.6092
8.437 | -90.6088
8.4373 | -90.6086
8.4376 | -90.6079
8.4379 | -90.6074
8.4382 | -90.6072
8.4386 | -90.6068
8.4389 | -90.606
8.4392 | -90.6055
8.4395 | -90.6054
8.4398 | -90.605
8.4401 | -90.6053
8.4404 | -90.6052
8.4407 | -90.6048
8.441 | -90.6045
8.4413 | -90.6042
8.4417 | -90.6038
8.442 | -90.6034
8.4423 | -90.602
8.4426 | -90.6015
8.4429 | -90.6018
8.4432 | -90.6014
8.4435 | -90.601
8.4438 | -90.6009
8.4441 | -90.6006
8.4445 | -90.5994
8.4448 | -90.5983
8.4451 | -90.5981
8.4454 | -90.5979
8.4457 | -90.5973
8.446 | -90.5969
8.4463 | -90.5964
8.4466 | -90.5959
8.4469 | -90.5958
8.4472 | -90.5957
8.4476 | -90.5954
8.4479 | -90.5949
8.4482 | -90.5947
8.4485 | -90.5944
8.4488 | -90.5944
8.4491 | -90.594
8.4494 | -90.594
8.4497 | -90.5944
8.45 | -90.5948
8.4504 | -90.5945
8.4507 | -90.5942
8.451 | -90.594
8.4513 | -90.5936
8.4516 | -90.5929
8.4519 | -90.5927
8.4522 | -90.5925
8.4525 | -90.5923
8.4528 | -90.5925
8.4531 | -90.593
8.4535 | -90.5927
8.4538 | -90.5919
8.4541 | -90.5917
8.4544 | -90.5912
8.4547 | -90.591
8.455 | -90.5909
8.4553 | -90.591
8.4556 | -90.5914
8.4559 | -90.5913
8.4563 | -90.5911
8.4566 | -90.5907
8.4569 | -90.5903
8.4572 | -90.59
8.4575 | -90.5892
8.4578 | -90.5892
8.4581 | -90.5895
8.4584 | -90.5894
8.4587 | -90.5887
8.459 | -90.588
8.4594 | -90.5879
8.4597 | -90.5883
8.46 | -90.588
8.4603 | -90.5873
8.4606 | -90.5867
8.4609 | -90.5859
8.4612 | -90.585
8.4615 | -90.5843
8.4618 | -90.5839
8.4622 | -90.5838
8.4625 | -90.5834
8.4628 | -90.5825
8.4631 | -90.5823
8.4634 | -90.5819
8.4637 | -90.5815
8.464 | -90.5814
8.4643 | -90.5812
8.4646 | -90.5812
8.4649 | -90.5805
8.4653 | -90.5801
8.4656 | -90.5802
8.4659 | -90.58
8.4662 | -90.5803
8.4665 | -90.58
8.4668 | -90.5796
8.4671 | -90.5797
8.4674 | -90.5798
8.4677 | -90.5798
8.4681 | -90.5803
8.4684 | -90.5806
8.4687 | -90.5806
8.469 | -90.5803
8.4693 | -90.5802
8.4696 | -90.58
8.4699 | -90.58
8.4702 | -90.5799
8.4705 | -90.5796
8.4709 | -90.5795
8.4712 | -90.5794
8.4715 | -90.5791
8.4718 | -90.5789
8.4721 | -90.5787
8.4724 | -90.5786
8.4727 | -90.5785
8.473 | -90.5782
8.4733 | -90.5779
8.4736 | -90.5777
8.474 | -90.5773
8.4743 | -90.5773
8.4746 | -90.5775
8.4749 | -90.578
8.4752 | -90.5784
8.4755 | -90.5784
8.4758 | -90.5781
8.4761 | -90.5775
8.4764 | -90.5776
8.4767 | -90.5783
8.4771 | -90.5786
8.4774 | -90.5784
8.4777 | -90.5783
8.478 | -90.5783
8.4783 | -90.5782
8.4786 | -90.5775
8.4789 | -90.5771
8.4792 | -90.5768
8.4795 | -90.5758
8.4799 | -90.5747
8.4802 | -90.5746
8.4805 | -90.5745
8.4808 | -90.5743
8.4811 | -90.5741
8.4814 | -90.5739
8.4817 | -90.5737
8.482 | -90.5732
8.4823 | -90.5725
8.4826 | -90.5723
8.483 | -90.5724
8.4833 | -90.5729
8.4836 | -90.5729
8.4839 | -90.5724
8.4842 | -90.5726
8.4845 | -90.5729
8.4848 | -90.5726
8.4851 | -90.5733
8.4854 | -90.573
8.4858 | -90.5731
8.4861 | -90.5725
8.4864 | -90.5722
8.4867 | -90.5721
8.487 | -90.5719
8.4873 | -90.5716
8.4876 | -90.5715
8.4879 | -90.5711
8.4882 | -90.5709
8.4886 | -90.5712
8.4889 | -90.5712
8.4892 | -90.5714
8.4895 | -90.5723
8.4898 | -90.5726
8.4901 | -90.5727
8.4904 | -90.5727
8.4907 | -90.5727
8.491 | -90.5726
8.4913 | -90.5723
8.4917 | -90.572
8.492 | -90.5719
8.4923 | -90.5716
8.4926 | -90.5715
8.4929 | -90.5718
8.4932 | -90.5714
8.4935 | -90.5711
8.4938 | -90.5716
8.4941 | -90.5714
8.4945 | -90.5716
8.4948 | -90.5717
8.4951 | -90.5717
8.4954 | -90.5717
8.4957 | -90.5714
8.496 | -90.5712
8.4963 | -90.5715
8.4966 | -90.5715
8.4969 | -90.5715
8.4972 | -90.5718
8.4976 | -90.5721
8.4979 | -90.5719
8.4982 | -90.5719
8.4985 | -90.5722
8.4988 | -90.5723
8.4991 | -90.5728
8.4994 | -90.5732
8.4997 | -90.5733
8.5 | -90.5733
8.5004 | -90.5735
8.5007 | -90.5735
8.501 | -90.5736
8.5013 | -90.5736
8.5016 | -90.5734
8.5019 | -90.5733
8.5022 | -90.5735
8.5025 | -90.5738
8.5028 | -90.574
8.5031 | -90.5743
8.5035 | -90.5748
8.5038 | -90.575
8.5041 | -90.5749
8.5044 | -90.5749
8.5047 | -90.5749
8.505 | -90.5751
8.5053 | -90.5751
8.5056 | -90.5751
8.5059 | -90.5754
8.5063 | -90.5756
8.5066 | -90.5763
8.5069 | -90.5769
8.5072 | -90.5771
8.5075 | -90.5771
8.5078 | -90.5773
8.5081 | -90.5775
8.5084 | -90.5775
8.5087 | -90.5775
8.509 | -90.5776
8.5094 | -90.5778
8.5097 | -90.5776
8.51 | -90.5778
8.5103 | -90.5778
8.5106 | -90.5778
8.5109 | -90.5776
8.5112 | -90.5774
8.5115 | -90.5781
8.5118 | -90.5782
8.5122 | -90.5783
8.5125 | -90.5783
8.5128 | -90.5783
8.5131 | -90.5784
8.5134 | -90.5785
8.5137 | -90.5787
8.514 | -90.5789
8.5143 | -90.579
8.5146 | -90.5789
8.5149 | -90.5792
8.5153 | -90.58
8.5156 | -90.5803
8.5159 | -90.5806
8.5162 | -90.5808
8.5165 | -90.581
8.5168 | -90.5812
8.5171 | -90.5814
8.5174 | -90.5816
8.5177 | -90.5818
8.5181 | -90.5819
8.5184 | -90.5818
8.5187 | -90.5812
8.519 | -90.5812
8.5193 | -90.5823
8.5196 | -90.5823
8.5199 | -90.582
8.5202 | -90.5822
8.5205 | -90.5822
8.5208 | -90.5821
8.5212 | -90.5825
8.5215 | -90.5831
8.5218 | -90.5834
8.5221 | -90.5836
8.5224 | -90.5841
8.5227 | -90.5844
8.523 | -90.5849
8.5233 | -90.5854
8.5236 | -90.5858
8.524 | -90.5859
8.5243 | -90.586
8.5246 | -90.5864
8.5249 | -90.5866
8.5252 | -90.5871
8.5255 | -90.5877
8.5258 | -90.588
8.5261 | -90.588
8.5264 | -90.5877
8.5267 | -90.5874
8.5271 | -90.5866
8.5274 | -90.5866
8.5277 | -90.5868
8.528 | -90.5872
8.5283 | -90.5874
8.5286 | -90.5876
8.5289 | -90.5878
8.5292 | -90.5881
8.5295 | -90.5884
8.5299 | -90.589
8.5302 | -90.5896
8.5305 | -90.5898
8.5308 | -90.5893
8.5311 | -90.5894
8.5314 | -90.5897
8.5317 | -90.59
8.532 | -90.5903
8.5323 | -90.591
8.5326 | -90.5915
8.533 | -90.5916
8.5333 | -90.5919
8.5336 | -90.5922
8.5339 | -90.5925
8.5342 | -90.5932
8.5345 | -90.5938
8.5348 | -90.5944
8.5351 | -90.5946
8.5354 | -90.5947
8.5358 | -90.5948
8.5361 | -90.595
8.5364 | -90.5951
8.5367 | -90.5952
8.537 | -90.5957
8.5373 | -90.5963
8.5376 | -90.597
8.5379 | -90.5983
8.5382 | -90.5985
8.5385 | -90.5995
8.5389 | -90.5998
8.5392 | -90.6
8.5395 | -90.5999
8.5398 | -90.6005
8.5401 | -90.6012
8.5404 | -90.6016
8.5407 | -90.602
8.541 | -90.6024
8.5413 | -90.6025
8.5417 | -90.6028
8.542 | -90.6033
8.5423 | -90.6035
8.5426 | -90.6038
8.5429 | -90.6042
8.5432 | -90.6047
8.5435 | -90.6053
8.5438 | -90.6061
8.5441 | -90.6067
8.5444 | -90.6072
8.5448 | -90.6087
8.5451 | -90.61
8.5454 | -90.6099
8.5457 | -90.61
8.546 | -90.6104
8.5463 | -90.6103
8.5466 | -90.6104
8.5469 | -90.6109
8.5472 | -90.6113
8.5476 | -90.6118
8.5479 | -90.6123
8.5482 | -90.6121
8.5485 | -90.6124
8.5488 | -90.6132
8.5491 | -90.6134
8.5494 | -90.6136
8.5497 | -90.6138
8.55 | -90.6144
8.5503 | -90.6158
8.5507 | -90.6164
8.551 | -90.617
8.5513 | -90.6173
8.5516 | -90.6177
8.5519 | -90.6182
8.5522 | -90.6188
8.5525 | -90.6189
8.5528 | -90.6191
8.5531 | -90.6197
8.5535 | -90.6203
8.5538 | -90.6207
8.5541 | -90.621
8.5544 | -90.6212
8.5547 | -90.6213
8.555 | -90.6214
8.5553 | -90.6218
8.5556 | -90.6224
8.5559 | -90.6227
8.5562 | -90.6224
8.5566 | -90.6231
8.5569 | -90.6236
8.5572 | -90.6242
8.5575 | -90.6248
8.5578 | -90.6252
8.5581 | -90.6257
8.5584 | -90.626
8.5587 | -90.6265
8.559 | -90.627
8.5594 | -90.6275
8.5597 | -90.6279
8.56 | -90.6285
8.5603 | -90.629
8.5606 | -90.6295
8.5609 | -90.6302
8.5612 | -90.6309
8.5615 | -90.6317
8.5618 | -90.6328
8.5621 | -90.6338
8.5625 | -90.6341
8.5628 | -90.6345
8.5631 | -90.6352
8.5634 | -90.6364
8.5637 | -90.6375
8.564 | -90.638
8.5643 | -90.6386
8.5646 | -90.6392
8.5649 | -90.6398
8.5653 | -90.6412
8.5656 | -90.6415
8.5659 | -90.6419
8.5662 | -90.6426
8.5665 | -90.6428
8.5668 | -90.643
8.5671 | -90.6433
8.5674 | -90.6436
8.5677 | -90.6441
8.568 | -90.6447
8.5684 | -90.6456
8.5687 | -90.6461
8.569 | -90.6465
8.5693 | -90.6469
8.5696 | -90.6474
8.5699 | -90.6481
8.5702 | -90.6488
8.5705 | -90.6493
8.5708 | -90.6495
8.5712 | -90.6499
8.5715 | -90.6505
8.5718 | -90.6512
8.5721 | -90.6518
8.5724 | -90.6523
8.5727 | -90.6528
8.573 | -90.6532
8.5733 | -90.6536
8.5736 | -90.6543
8.5739 | -90.6552
8.5743 | -90.6564
8.5746 | -90.6571
8.5749 | -90.658
8.5752 | -90.6594
8.5755 | -90.6602
8.5758 | -90.6608
8.5761 | -90.6614
8.5764 | -90.6622
8.5767 | -90.6629
8.5771 | -90.6632
8.5774 | -90.6638
8.5777 | -90.6646
8.578 | -90.6651
8.5783 | -90.6656
8.5786 | -90.6664
8.5789 | -90.667
8.5792 | -90.6677
8.5795 | -90.6684
8.5798 | -90.6687
8.5802 | -90.6692
8.5805 | -90.6701
8.5808 | -90.671
8.5811 | -90.6715
8.5814 | -90.672
8.5817 | -90.6727
8.582 | -90.6736
8.5823 | -90.6743
8.5826 | -90.6752
8.583 | -90.676
8.5833 | -90.6773
8.5836 | -90.6777
8.5839 | -90.6786
8.5842 | -90.6794
8.5845 | -90.6802
8.5848 | -90.6811
8.5851 | -90.6817
8.5854 | -90.6827
8.5857 | -90.6832
8.5861 | -90.6839
8.5864 | -90.685
8.5867 | -90.6859
8.587 | -90.6865
8.5873 | -90.6872
8.5876 | -90.6883
8.5879 | -90.6891
8.5882 | -90.6898
8.5885 | -90.6901
8.5889 | -90.6908
8.5892 | -90.6915
8.5895 | -90.6924
8.5898 | -90.6933
8.5901 | -90.6942
8.5904 | -90.6951
8.5907 | -90.6955
8.591 | -90.6959
8.5913 | -90.6971
8.5916 | -90.6981
8.592 | -90.6992
8.5923 | -90.7006
8.5926 | -90.7016
8.5929 | -90.7024
8.5932 | -90.7029
8.5935 | -90.7034
8.5938 | -90.7038
8.5941 | -90.7046
8.5944 | -90.7052
8.5948 | -90.7057
8.5951 | -90.7058
8.5954 | -90.706
8.5957 | -90.7066
8.596 | -90.7069
8.5963 | -90.7073
8.5966 | -90.708
8.5969 | -90.7091
8.5972 | -90.7101
8.5975 | -90.7108
8.5979 | -90.7116
8.5982 | -90.7124
8.5985 | -90.7131
8.5988 | -90.7138
8.5991 | -90.7141
8.5994 | -90.7143
8.5997 | -90.7152
8.6 | -90.7163
8.6003 | -90.7174
8.6007 | -90.7181
8.601 | -90.719
8.6013 | -90.72
8.6016 | -90.7202
8.6019 | -90.7207
8.6022 | -90.7213
8.6025 | -90.7219
8.6028 | -90.723
8.6031 | -90.724
8.6034 | -90.725
8.6038 | -90.7261
8.6041 | -90.7268
8.6044 | -90.7275
8.6047 | -90.7285
8.605 | -90.7303
8.6053 | -90.7321
8.6056 | -90.7332
8.6059 | -90.7344
8.6062 | -90.7353
8.6066 | -90.7357
8.6069 | -90.7363
8.6072 | -90.7372
8.6075 | -90.7383
8.6078 | -90.7394
8.6081 | -90.7402
8.6084 | -90.7411
8.6087 | -90.742
8.609 | -90.7433
8.6093 | -90.7444
8.6097 | -90.7448
8.61 | -90.7458
8.6103 | -90.7474
8.6106 | -90.7485
8.6109 | -90.749
8.6112 | -90.7501
8.6115 | -90.7516
8.6118 | -90.7521
8.6121 | -90.753
8.6125 | -90.7543
8.6128 | -90.7561
8.6131 | -90.7579
8.6134 | -90.7588
8.6137 | -90.7595
8.614 | -90.76
8.6143 | -90.7607
8.6146 | -90.7615
8.6149 | -90.7623
8.6152 | -90.7634
8.6156 | -90.7644
8.6159 | -90.7649
8.6162 | -90.7656
8.6165 | -90.7668
8.6168 | -90.7679
8.6171 | -90.7689
8.6174 | -90.7696
8.6177 | -90.7705
8.618 | -90.7711
8.6184 | -90.7724
8.6187 | -90.7733
8.619 | -90.774
8.6193 | -90.7756
8.6196 | -90.7765
8.6199 | -90.7772
8.6202 | -90.7781
8.6205 | -90.7793
8.6208 | -90.7802
8.6212 | -90.7813
8.6215 | -90.7823
8.6218 | -90.7825
8.6221 | -90.7829
8.6224 | -90.7839
8.6227 | -90.7847
8.623 | -90.7854
8.6233 | -90.7866
8.6236 | -90.7876
8.6239 | -90.7884
8.6243 | -90.7892
8.6246 | -90.7905
8.6249 | -90.7918
8.6252 | -90.7926
8.6255 | -90.7935
8.6258 | -90.7946
8.6261 | -90.7956
8.6264 | -90.7967
8.6267 | -90.798
8.6271 | -90.7989
8.6274 | -90.7999
8.6277 | -90.8009
8.628 | -90.802
8.6283 | -90.8025
8.6286 | -90.8032
8.6289 | -90.8042
8.6292 | -90.8051
8.6295 | -90.8058
8.6298 | -90.8068
8.6302 | -90.8078
8.6305 | -90.8088
8.6308 | -90.8097
8.6311 | -90.8108
8.6314 | -90.8117
8.6317 | -90.813
8.632 | -90.8142
8.6323 | -90.8146
8.6326 | -90.815
8.633 | -90.816
8.6333 | -90.8182
8.6336 | -90.8202
8.6339 | -90.8209
8.6342 | -90.8214
8.6345 | -90.8224
8.6348 | -90.8236
8.6351 | -90.8249
8.6354 | -90.8256
8.6357 | -90.8258
8.6361 | -90.8271
8.6364 | -90.8284
8.6367 | -90.8293
8.637 | -90.8308
8.6373 | -90.8321
8.6376 | -90.833
8.6379 | -90.8334
8.6382 | -90.8341
8.6385 | -90.8352
8.6389 | -90.8358
8.6392 | -90.8362
8.6395 | -90.8372
8.6398 | -90.8387
8.6401 | -90.8398
8.6404 | -90.8407
8.6407 | -90.8417
8.641 | -90.8428
8.6413 | -90.8438
8.6416 | -90.8447
8.642 | -90.8457
8.6423 | -90.8466
8.6426 | -90.8477
8.6429 | -90.8487
8.6432 | -90.8499
8.6435 | -90.8508
8.6438 | -90.8518
8.6441 | -90.8528
8.6444 | -90.8528
8.6448 | -90.8532
8.6451 | -90.8542
8.6454 | -90.8554
8.6457 | -90.8565
8.646 | -90.8579
8.6463 | -90.8586
8.6466 | -90.8593
8.6469 | -90.8604
8.6472 | -90.8616
8.6475 | -90.8627
8.6479 | -90.864
8.6482 | -90.8652
8.6485 | -90.8665
8.6488 | -90.8675
8.6491 | -90.8685
8.6494 | -90.8696
8.6497 | -90.8707
8.65 | -90.8717
8.6503 | -90.8734
8.6507 | -90.8753
8.651 | -90.8762
8.6513 | -90.877
8.6516 | -90.8782
8.6519 | -90.8794
8.6522 | -90.8805
8.6525 | -90.8811
8.6528 | -90.882
8.6531 | -90.8838
8.6534 | -90.8854
8.6538 | -90.8865
8.6541 | -90.8876
8.6544 | -90.8888
8.6547 | -90.8899
8.655 | -90.8909
8.6553 | -90.8919
8.6556 | -90.8934
8.6559 | -90.8945
8.6562 | -90.8955
8.6566 | -90.8964
8.6569 | -90.8975
8.6572 | -90.8988
8.6575 | -90.8997
8.6578 | -90.9009
8.6581 | -90.9022
8.6584 | -90.9033
8.6587 | -90.9051
8.659 | -90.9064
8.6593 | -90.9075
8.6597 | -90.9088
8.66 | -90.9102
8.6603 | -90.9113
8.6606 | -90.9127
8.6609 | -90.9138
8.6612 | -90.915
8.6615 | -90.9165
8.6618 | -90.918
8.6621 | -90.9194
8.6625 | -90.9206
8.6628 | -90.9218
8.6631 | -90.9228
8.6634 | -90.9236
8.6637 | -90.9246
8.664 | -90.9261
8.6643 | -90.9274
8.6646 | -90.9287
8.6649 | -90.9301
8.6652 | -90.9314
8.6656 | -90.9329
8.6659 | -90.934
8.6662 | -90.935
8.6665 | -90.9361
8.6668 | -90.9374
8.6671 | -90.9387
8.6674 | -90.9402
8.6677 | -90.9421
8.668 | -90.9439
8.6684 | -90.9453
8.6687 | -90.9465
8.669 | -90.9476
8.6693 | -90.9491
8.6696 | -90.95
8.6699 | -90.9509
8.6702 | -90.952
8.6705 | -90.9532
8.6708 | -90.9543
8.6711 | -90.955
8.6715 | -90.9557
8.6718 | -90.9563
8.6721 | -90.9573
8.6724 | -90.959
8.6727 | -90.9602
8.673 | -90.9618
8.6733 | -90.9627
8.6736 | -90.9639
8.6739 | -90.9661
8.6743 | -90.9676
8.6746 | -90.969
8.6749 | -90.9702
8.6752 | -90.9716
8.6755 | -90.9733
8.6758 | -90.9746
8.6761 | -90.9757
8.6764 | -90.977
8.6767 | -90.9778
8.677 | -90.9788
8.6774 | -90.9798
8.6777 | -90.9809
8.678 | -90.9821
8.6783 | -90.9828
8.6786 | -90.9841
8.6789 | -90.9863
8.6792 | -90.9875
8.6795 | -90.9892
8.6798 | -90.9907
8.6802 | -90.992
8.6805 | -90.993
8.6808 | -90.9937
8.6811 | -90.9948
8.6814 | -90.9959
8.6817 | -90.9971
8.682 | -90.9986
8.6823 | -90.9998
8.6826 | -91.0009
8.6829 | -91.0023
8.6833 | -91.0039
8.6836 | -91.0053
8.6839 | -91.0059
8.6842 | -91.0062
8.6845 | -91.0075
8.6848 | -91.0087
8.6851 | -91.0102
8.6854 | -91.0129
8.6857 | -91.0151
8.6861 | -91.0165
8.6864 | -91.0176
8.6867 | -91.0189
8.687 | -91.0205
8.6873 | -91.0219
8.6876 | -91.023
8.6879 | -91.0241
8.6882 | -91.0252
8.6885 | -91.0265
8.6888 | -91.0282
8.6892 | -91.0297
8.6895 | -91.0316
8.6898 | -91.0331
8.6901 | -91.0344
8.6904 | -91.0358
8.6907 | -91.0371
8.691 | -91.0383
8.6913 | -91.0398
8.6916 | -91.0413
8.692 | -91.0425
8.6923 | -91.0439
8.6926 | -91.0451
8.6929 | -91.0473
8.6932 | -91.0489
8.6935 | -91.0501
8.6938 | -91.0516
8.6941 | -91.0527
8.6944 | -91.054
8.6947 | -91.0554
8.6951 | -91.0567
8.6954 | -91.058
8.6957 | -91.0595
8.696 | -91.0606
8.6963 | -91.0618
8.6966 | -91.063
8.6969 | -91.0642
8.6972 | -91.0652
8.6975 | -91.066
8.6979 | -91.0671
8.6982 | -91.0684
8.6985 | -91.0699
8.6988 | -91.0711
8.6991 | -91.0715
8.6994 | -91.0722
8.6997 | -91.0735
8.7 | -91.0746
8.7003 | -91.0754
8.7006 | -91.0764
8.701 | -91.0775
8.7013 | -91.0785
8.7016 | -91.0797
8.7019 | -91.0809
8.7022 | -91.0821
8.7025 | -91.0834
8.7028 | -91.0846
8.7031 | -91.086
8.7034 | -91.0873
8.7038 | -91.0881
8.7041 | -91.0891
8.7044 | -91.0904
8.7047 | -91.0914
8.705 | -91.0926
8.7053 | -91.094
8.7056 | -91.0954
8.7059 | -91.0969
8.7062 | -91.0981
8.7065 | -91.0994
8.7069 | -91.1003
8.7072 | -91.1011
8.7075 | -91.1025
8.7078 | -91.1041
8.7081 | -91.1052
8.7084 | -91.1062
8.7087 | -91.1068
8.709 | -91.1081
8.7093 | -91.1093
8.7097 | -91.1104
8.71 | -91.1116
8.7103 | -91.1128
8.7106 | -91.1142
8.7109 | -91.1154
8.7112 | -91.1165
8.7115 | -91.1174
8.7118 | -91.1187
8.7121 | -91.1198
8.7124 | -91.1209
8.7128 | -91.1221
8.7131 | -91.1236
8.7134 | -91.1256
8.7137 | -91.1269
8.714 | -91.1289
8.7143 | -91.1305
8.7146 | -91.1318
8.7149 | -91.1333
8.7152 | -91.1354
8.7156 | -91.1366
8.7159 | -91.1378
8.7162 | -91.1389
8.7165 | -91.141
8.7168 | -91.1422
8.7171 | -91.1437
8.7174 | -91.1452
8.7177 | -91.146
8.718 | -91.1474
8.7183 | -91.1493
8.7187 | -91.1513
8.719 | -91.1532
8.7193 | -91.1545
8.7196 | -91.1564
8.7199 | -91.1602
8.7202 | -91.1617
8.7205 | -91.1641
8.7208 | -91.1669
8.7211 | -91.1689
8.7215 | -91.1705
8.7218 | -91.1726
8.7221 | -91.1753
8.7224 | -91.1766
8.7227 | -91.179
8.723 | -91.1829
8.7233 | -91.1847
8.7236 | -91.1859
8.7239 | -91.1885
8.7242 | -91.1893
8.7246 | -91.1897
8.7249 | -91.1902
8.7252 | -91.1916
8.7255 | -91.1928
8.7258 | -91.1946
8.7261 | -91.1967
8.7264 | -91.1982
8.7267 | -91.198
8.727 | -91.1992
8.7274 | -91.1998
8.7277 | -91.201
8.728 | -91.2021
8.7283 | -91.2032
8.7286 | -91.2049
8.7289 | -91.206
8.7292 | -91.2071
8.7295 | -91.2089
8.7298 | -91.2109
8.7301 | -91.2121
8.7305 | -91.2134
8.7308 | -91.2141
8.7311 | -91.2154
8.7314 | -91.2166
8.7317 | -91.2179
8.732 | -91.219
8.7323 | -91.2205
8.7326 | -91.2216
8.7329 | -91.2233
8.7333 | -91.2241
8.7336 | -91.2251
8.7339 | -91.2263
8.7342 | -91.2275
8.7345 | -91.2289
8.7348 | -91.2301
8.7351 | -91.2308
8.7354 | -91.2329
8.7357 | -91.2343
8.736 | -91.2367
8.7364 | -91.2379
8.7367 | -91.2388
8.737 | -91.2405
8.7373 | -91.2424
8.7376 | -91.2437
8.7379 | -91.2452
8.7382 | -91.2467
8.7385 | -91.2477
8.7388 | -91.2489
8.7392 | -91.2502
8.7395 | -91.2512
8.7398 | -91.2534
8.7401 | -91.2555
8.7404 | -91.2565
8.7407 | -91.2576
8.741 | -91.2587
8.7413 | -91.2593
8.7416 | -91.2596
8.7419 | -91.2605
8.7423 | -91.2613
8.7426 | -91.2624
8.7429 | -91.2639
8.7432 | -91.2652
8.7435 | -91.2662
8.7438 | -91.2673
8.7441 | -91.2684
8.7444 | -91.2699
8.7447 | -91.271
8.7451 | -91.272
8.7454 | -91.2731
8.7457 | -91.2742
8.746 | -91.2751
8.7463 | -91.2762
8.7466 | -91.2773
8.7469 | -91.2782
8.7472 | -91.2793
8.7475 | -91.2804
8.7478 | -91.2813
8.7482 | -91.2822
8.7485 | -91.2837
8.7488 | -91.285
8.7491 | -91.2858
8.7494 | -91.2868
8.7497 | -91.2881
8.75 | -91.2896
8.7503 | -91.2908
8.7506 | -91.2919
8.751 | -91.2935
8.7513 | -91.2949
8.7516 | -91.296
8.7519 | -91.2971
8.7522 | -91.2982
8.7525 | -91.2994
8.7528 | -91.3008
8.7531 | -91.302
8.7534 | -91.3031
8.7538 | -91.304
8.7541 | -91.305
8.7544 | -91.3064
8.7547 | -91.3077
8.755 | -91.3088
8.7553 | -91.31
8.7556 | -91.3112
8.7559 | -91.3125
8.7562 | -91.3136
8.7565 | -91.3141
8.7569 | -91.3153
8.7572 | -91.3167
8.7575 | -91.3178
8.7578 | -91.3188
8.7581 | -91.3201
8.7584 | -91.3209
8.7587 | -91.3217
8.759 | -91.3226
8.7593 | -91.3237
8.7597 | -91.3256
8.76 | -91.3275
8.7603 | -91.3287
8.7606 | -91.3299
8.7609 | -91.3314
8.7612 | -91.3323
8.7615 | -91.3328
8.7618 | -91.3333
8.7621 | -91.3334
8.7624 | -91.3345
8.7628 | -91.3355
8.7631 | -91.3364
8.7634 | -91.3376
8.7637 | -91.3392
8.764 | -91.3409
8.7643 | -91.343
8.7646 | -91.3444
8.7649 | -91.3453
8.7652 | -91.3465
8.7655 | -91.3478
8.7659 | -91.3491
8.7662 | -91.3508
8.7665 | -91.3525
8.7668 | -91.3536
8.7671 | -91.3548
8.7674 | -91.3564
8.7677 | -91.3578
8.768 | -91.3589
8.7683 | -91.3599
8.7687 | -91.362
8.769 | -91.3638
8.7693 | -91.365
8.7696 | -91.3664
8.7699 | -91.3676
8.7702 | -91.3689
8.7705 | -91.3699
8.7708 | -91.3714
8.7711 | -91.3736
8.7714 | -91.3749
8.7718 | -91.3763
8.7721 | -91.378
8.7724 | -91.3799
8.7727 | -91.3813
8.773 | -91.3821
8.7733 | -91.3832
8.7736 | -91.3846
8.7739 | -91.3862
8.7742 | -91.388
8.7746 | -91.3887
8.7749 | -91.3893
8.7752 | -91.3905
8.7755 | -91.3918
8.7758 | -91.393
8.7761 | -91.394
8.7764 | -91.3951
8.7767 | -91.3958
8.777 | -91.3968
8.7774 | -91.398
8.7777 | -91.3984
8.778 | -91.3993
8.7783 | -91.4002
8.7786 | -91.4014
8.7789 | -91.4026
8.7792 | -91.404
8.7795 | -91.4054
8.7798 | -91.4065
8.7801 | -91.4074
8.7805 | -91.4079
8.7808 | -91.4088
8.7811 | -91.41
8.7814 | -91.4113
8.7817 | -91.412
8.782 | -91.4131
8.7823 | -91.4146
8.7826 | -91.416
8.7829 | -91.417
8.7833 | -91.4177
8.7836 | -91.4189
8.7839 | -91.4201
8.7842 | -91.4217
8.7845 | -91.4229
8.7848 | -91.4237
8.7851 | -91.4248
8.7854 | -91.4263
8.7857 | -91.4285
8.786 | -91.43
8.7864 | -91.4308
8.7867 | -91.4313
8.787 | -91.4323
8.7873 | -91.4333
8.7876 | -91.4344
8.7879 | -91.4354
8.7882 | -91.4364
8.7885 | -91.4375
8.7888 | -91.4385
8.7892 | -91.4397
8.7895 | -91.4409
8.7898 | -91.4421
8.7901 | -91.4432
8.7904 | -91.4445
8.7907 | -91.4457
8.791 | -91.4469
8.7913 | -91.4479
8.7916 | -91.4489
8.7919 | -91.4505
8.7923 | -91.4519
8.7926 | -91.453
8.7929 | -91.4541
8.7932 | -91.4554
8.7935 | -91.4566
8.7938 | -91.4577
8.7941 | -91.4587
8.7944 | -91.46
8.7947 | -91.4608
8.7951 | -91.4617
8.7954 | -91.4628
8.7957 | -91.4642
8.796 | -91.4659
8.7963 | -91.4672
8.7966 | -91.4688
8.7969 | -91.4705
8.7972 | -91.4715
8.7975 | -91.4724
8.7978 | -91.4735
8.7982 | -91.4748
8.7985 | -91.476
8.7988 | -91.4778
8.7991 | -91.4794
8.7994 | -91.4805
8.7997 | -91.4819
8.8 | -91.4843
8.8003 | -91.486
8.8006 | -91.4872
8.801 | -91.4883
8.8013 | -91.4891
8.8016 | -91.4905
8.8019 | -91.4916
8.8022 | -91.4932
8.8025 | -91.4939
8.8028 | -91.4941
8.8031 | -91.4946
8.8034 | -91.4955
8.8037 | -91.497
8.8041 | -91.4984
8.8044 | -91.4993
8.8047 | -91.4999
8.805 | -91.4999
8.8053 | -91.5006
8.8056 | -91.5017
8.8059 | -91.5026
8.8062 | -91.5035
8.8065 | -91.5045
8.8069 | -91.5057
8.8072 | -91.5061
8.8075 | -91.5068
8.8078 | -91.5073
8.8081 | -91.5071
8.8084 | -91.5078
8.8087 | -91.5101
8.809 | -91.5115
8.8093 | -91.5128
8.8096 | -91.5137
8.81 | -91.5147
8.8103 | -91.5155
8.8106 | -91.5167
8.8109 | -91.5178
8.8112 | -91.519
8.8115 | -91.5204
8.8118 | -91.5216
8.8121 | -91.5229
8.8124 | -91.5243
8.8128 | -91.5254
8.8131 | -91.5264
8.8134 | -91.5272
8.8137 | -91.528
8.814 | -91.5291
8.8143 | -91.5301
8.8146 | -91.5308
8.8149 | -91.5321
8.8152 | -91.5335
8.8155 | -91.5344
8.8159 | -91.5355
8.8162 | -91.5367
8.8165 | -91.5377
8.8168 | -91.539
8.8171 | -91.5409
8.8174 | -91.5431
8.8177 | -91.544
8.818 | -91.5452
8.8183 | -91.5471
8.8187 | -91.5487
8.819 | -91.55
8.8193 | -91.551
8.8196 | -91.5528
8.8199 | -91.5551
8.8202 | -91.5562
8.8205 | -91.5572
8.8208 | -91.5586
8.8211 | -91.5602
8.8214 | -91.5612
8.8218 | -91.5622
8.8221 | -91.5631
8.8224 | -91.564
8.8227 | -91.5649
8.823 | -91.5662
8.8233 | -91.5672
8.8236 | -91.5676
8.8239 | -91.5688
8.8242 | -91.5696
8.8246 | -91.5706
8.8249 | -91.572
8.8252 | -91.5724
8.8255 | -91.5731
8.8258 | -91.574
8.8261 | -91.5759
8.8264 | -91.5774
8.8267 | -91.5782
8.827 | -91.5791
8.8273 | -91.5799
8.8277 | -91.5812
8.828 | -91.5825
8.8283 | -91.5835
8.8286 | -91.5846
8.8289 | -91.5859
8.8292 | -91.5864
8.8295 | -91.5872
8.8298 | -91.5886
8.8301 | -91.5893
8.8305 | -91.5901
8.8308 | -91.5911
8.8311 | -91.5923
8.8314 | -91.5931
8.8317 | -91.5938
8.832 | -91.5951
8.8323 | -91.5965
8.8326 | -91.5977
8.8329 | -91.5988
8.8332 | -91.6002
8.8336 | -91.6012
8.8339 | -91.6023
8.8342 | -91.6034
8.8345 | -91.6044
8.8348 | -91.6054
8.8351 | -91.6064
8.8354 | -91.6075
8.8357 | -91.6086
8.836 | -91.6101
8.8364 | -91.6116
8.8367 | -91.6127
8.837 | -91.6136
8.8373 | -91.6146
8.8376 | -91.6156
8.8379 | -91.6163
8.8382 | -91.6172
8.8385 | -91.6182
8.8388 | -91.6193
8.8391 | -91.6204
8.8395 | -91.6216
8.8398 | -91.6237
8.8401 | -91.6256
8.8404 | -91.6273
8.8407 | -91.6286
8.841 | -91.6293
8.8413 | -91.6303
8.8416 | -91.6313
8.8419 | -91.6319
8.8423 | -91.633
8.8426 | -91.6338
8.8429 | -91.6345
8.8432 | -91.6339
8.8435 | -91.6347
8.8438 | -91.6357
8.8441 | -91.6368
8.8444 | -91.6378
8.8447 | -91.6391
8.845 | -91.641
8.8454 | -91.6424
8.8457 | -91.6436
8.846 | -91.6456
8.8463 | -91.6476
8.8466 | -91.6493
8.8469 | -91.6502
8.8472 | -91.6512
8.8475 | -91.6519
8.8478 | -91.6527
8.8482 | -91.6537
8.8485 | -91.6547
8.8488 | -91.6558
8.8491 | -91.6568
8.8494 | -91.6577
8.8497 | -91.6588
8.85 | -91.6598
8.8503 | -91.6608
8.8506 | -91.6612
8.8509 | -91.6619
8.8513 | -91.6634
8.8516 | -91.6643
8.8519 | -91.6647
8.8522 | -91.6655
8.8525 | -91.6665
8.8528 | -91.6675
8.8531 | -91.6685
8.8534 | -91.6687
8.8537 | -91.6691
8.8541 | -91.6701
8.8544 | -91.6713
8.8547 | -91.6724
8.855 | -91.6735
8.8553 | -91.6746
8.8556 | -91.6765
8.8559 | -91.678
8.8562 | -91.6793
8.8565 | -91.6805
8.8568 | -91.6814
8.8572 | -91.6826
8.8575 | -91.6838
8.8578 | -91.6848
8.8581 | -91.6859
8.8584 | -91.6875
8.8587 | -91.6892
8.859 | -91.6907
8.8593 | -91.692
8.8596 | -91.6929
8.86 | -91.6936
8.8603 | -91.6948
8.8606 | -91.6963
8.8609 | -91.6972
8.8612 | -91.6982
8.8615 | -91.6998
8.8618 | -91.7017
8.8621 | -91.7032
8.8624 | -91.704
8.8627 | -91.7047
8.8631 | -91.7053
8.8634 | -91.7059
8.8637 | -91.7068
8.864 | -91.7078
8.8643 | -91.7088
8.8646 | -91.7095
8.8649 | -91.7104
8.8652 | -91.7118
8.8655 | -91.713
8.8659 | -91.7142
8.8662 | -91.7155
8.8665 | -91.7164
8.8668 | -91.7174
8.8671 | -91.7189
8.8674 | -91.7199
8.8677 | -91.7204
8.868 | -91.7211
8.8683 | -91.7221
8.8686 | -91.723
8.869 | -91.7246
8.8693 | -91.7258
8.8696 | -91.7268
8.8699 | -91.7273
8.8702 | -91.7283
8.8705 | -91.7287
8.8708 | -91.7293
8.8711 | -91.7309
8.8714 | -91.7323
8.8718 | -91.733
8.8721 | -91.7343
8.8724 | -91.7355
8.8727 | -91.7363
8.873 | -91.7375
8.8733 | -91.7387
8.8736 | -91.7396
8.8739 | -91.7408
8.8742 | -91.7423
8.8745 | -91.7429
8.8749 | -91.744
8.8752 | -91.7452
8.8755 | -91.7472
8.8758 | -91.7484
8.8761 | -91.7495
8.8764 | -91.7514
8.8767 | -91.7528
8.877 | -91.7537
8.8773 | -91.7547
8.8777 | -91.7557
8.878 | -91.7566
8.8783 | -91.7582
8.8786 | -91.759
8.8789 | -91.7599
8.8792 | -91.7608
8.8795 | -91.7617
8.8798 | -91.7628
8.8801 | -91.7639
8.8804 | -91.7645
8.8808 | -91.7651
8.8811 | -91.7663
8.8814 | -91.7675
8.8817 | -91.7684
8.882 | -91.7698
8.8823 | -91.7711
8.8826 | -91.7722
8.8829 | -91.7737
8.8832 | -91.7753
8.8836 | -91.7765
8.8839 | -91.7774
8.8842 | -91.7785
8.8845 | -91.78
8.8848 | -91.7812
8.8851 | -91.7815
8.8854 | -91.7829
8.8857 | -91.7842
8.886 | -91.7854
8.8863 | -91.7864
8.8867 | -91.7874
8.887 | -91.7893
8.8873 | -91.7908
8.8876 | -91.7918
8.8879 | -91.7928
8.8882 | -91.7934
8.8885 | -91.7941
8.8888 | -91.7947
8.8891 | -91.7962
8.8895 | -91.7988
8.8898 | -91.8004
8.8901 | -91.8015
8.8904 | -91.8025
8.8907 | -91.803
8.891 | -91.804
8.8913 | -91.8048
8.8916 | -91.8053
8.8919 | -91.8056
8.8923 | -91.8072
8.8926 | -91.8085
8.8929 | -91.8088
8.8932 | -91.8095
8.8935 | -91.8102
8.8938 | -91.8109
8.8941 | -91.8117
8.8944 | -91.8137
8.8947 | -91.8148
8.895 | -91.8152
8.8954 | -91.8157
8.8957 | -91.8163
8.896 | -91.817
8.8963 | -91.818
8.8966 | -91.8189
8.8969 | -91.8201
8.8972 | -91.8209
8.8975 | -91.8215
8.8978 | -91.8225
8.8981 | -91.8236
8.8985 | -91.8241
8.8988 | -91.8245
8.8991 | -91.8254
8.8994 | -91.8265
8.8997 | -91.8272
8.9 | -91.8281
8.9003 | -91.8285
8.9006 | -91.8291
8.9009 | -91.8295
8.9013 | -91.8301
8.9016 | -91.8309
8.9019 | -91.8323
8.9022 | -91.8334
8.9025 | -91.8346
8.9028 | -91.8367
8.9031 | -91.8389
8.9034 | -91.8402
8.9037 | -91.8411
8.904 | -91.8419
8.9044 | -91.8421
8.9047 | -91.8426
8.905 | -91.8435
8.9053 | -91.8443
8.9056 | -91.845
8.9059 | -91.8455
8.9062 | -91.8459
8.9065 | -91.8467
8.9068 | -91.8475
8.9072 | -91.8484
8.9075 | -91.8497
8.9078 | -91.8511
8.9081 | -91.8525
8.9084 | -91.8553
8.9087 | -91.8568
8.909 | -91.8578
8.9093 | -91.859
8.9096 | -91.8596
8.91 | -91.8601
8.9103 | -91.8618
8.9106 | -91.8626
8.9109 | -91.8632
8.9112 | -91.8648
8.9115 | -91.8679
8.9118 | -91.8684
8.9121 | -91.8693
8.9124 | -91.8705
8.9127 | -91.8719
8.9131 | -91.8729
8.9134 | -91.8736
8.9137 | -91.8745
8.914 | -91.8752
8.9143 | -91.8758
8.9146 | -91.8764
8.9149 | -91.8771
8.9152 | -91.8777
8.9155 | -91.8785
8.9159 | -91.8796
8.9162 | -91.8808
8.9165 | -91.8818
8.9168 | -91.8828
8.9171 | -91.884
8.9174 | -91.8855
8.9177 | -91.887
8.918 | -91.8879
8.9183 | -91.8889
8.9186 | -91.8899
8.919 | -91.8908
8.9193 | -91.8917
8.9196 | -91.8928
8.9199 | -91.894
8.9202 | -91.8946
8.9205 | -91.8959
8.9208 | -91.8965
8.9211 | -91.897
8.9214 | -91.8978
8.9218 | -91.8985
8.9221 | -91.8988
8.9224 | -91.8991
8.9227 | -91.8995
8.923 | -91.9003
8.9233 | -91.9014
8.9236 | -91.9023
8.9239 | -91.9033
8.9242 | -91.9041
8.9245 | -91.9047
8.9249 | -91.9055
8.9252 | -91.9063
8.9255 | -91.9071
8.9258 | -91.908
8.9261 | -91.9084
8.9264 | -91.9092
8.9267 | -91.9101
8.927 | -91.9109
8.9273 | -91.9114
8.9277 | -91.9126
8.928 | -91.9135
8.9283 | -91.9147
8.9286 | -91.9155
8.9289 | -91.9163
8.9292 | -91.9167
8.9295 | -91.9169
8.9298 | -91.9174
8.9301 | -91.919
8.9304 | -91.9207
8.9308 | -91.9218
8.9311 | -91.9225
8.9314 | -91.923
8.9317 | -91.9237
8.932 | -91.9243
8.9323 | -91.9254
8.9326 | -91.9264
8.9329 | -91.9269
8.9332 | -91.9275
8.9336 | -91.9284
8.9339 | -91.9291
8.9342 | -91.9299
8.9345 | -91.9309
8.9348 | -91.9318
8.9351 | -91.9328
8.9354 | -91.9332
8.9357 | -91.9334
8.936 | -91.9345
8.9363 | -91.9358
8.9367 | -91.9365
8.937 | -91.9374
8.9373 | -91.9392
8.9376 | -91.9405
8.9379 | -91.941
8.9382 | -91.9419
8.9385 | -91.9428
8.9388 | -91.9435
8.9391 | -91.9442
8.9395 | -91.9451
8.9398 | -91.946
8.9401 | -91.9469
8.9404 | -91.9478
8.9407 | -91.9488
8.941 | -91.9497
8.9413 | -91.9513
8.9416 | -91.9521
8.9419 | -91.9543
8.9422 | -91.9566
8.9426 | -91.9573
8.9429 | -91.9582
8.9432 | -91.9588
8.9435 | -91.9591
8.9438 | -91.9595
8.9441 | -91.9604
8.9444 | -91.9614
8.9447 | -91.9624
8.945 | -91.9637
8.9454 | -91.9645
8.9457 | -91.9651
8.946 | -91.9658
8.9463 | -91.9668
8.9466 | -91.9675
8.9469 | -91.9684
8.9472 | -91.9696
8.9475 | -91.9702
8.9478 | -91.9711
8.9481 | -91.9723
8.9485 | -91.9735
8.9488 | -91.9745
8.9491 | -91.9755
8.9494 | -91.9781
8.9497 | -91.9794
8.95 | -91.9805
8.9503 | -91.9818
8.9506 | -91.9815
8.9509 | -91.9827
8.9513 | -91.9839
8.9516 | -91.9846
8.9519 | -91.9854
8.9522 | -91.9861
8.9525 | -91.9867
8.9528 | -91.9872
8.9531 | -91.9879
8.9534 | -91.9885
8.9537 | -91.9891
8.954 | -91.9898
8.9544 | -91.9909
8.9547 | -91.9918
8.955 | -91.9926
8.9553 | -91.9934
8.9556 | -91.9942
8.9559 | -91.9947
8.9562 | -91.9954
8.9565 | -91.9959
8.9568 | -91.9961
8.9572 | -91.9969
8.9575 | -91.998
8.9578 | -91.9991
8.9581 | -91.9998
8.9584 | -92.0003
8.9587 | -92.0006
8.959 | -92.0018
8.9593 | -92.0032
8.9596 | -92.0041
8.9599 | -92.005
8.9603 | -92.0061
8.9606 | -92.0072
8.9609 | -92.0082
8.9612 | -92.0092
8.9615 | -92.0101
8.9618 | -92.0111
8.9621 | -92.013
8.9624 | -92.0143
8.9627 | -92.015
8.9631 | -92.0157
8.9634 | -92.0161
8.9637 | -92.0166
8.964 | -92.0172
8.9643 | -92.0178
8.9646 | -92.018
8.9649 | -92.0182
8.9652 | -92.0187
8.9655 | -92.0202
8.9658 | -92.0212
8.9662 | -92.0224
8.9665 | -92.0237
8.9668 | -92.0264
8.9671 | -92.0283
8.9674 | -92.0294
8.9677 | -92.0305
8.968 | -92.032
8.9683 | -92.0331
8.9686 | -92.0332
8.969 | -92.034
8.9693 | -92.0345
8.9696 | -92.0356
8.9699 | -92.0364
8.9702 | -92.037
8.9705 | -92.0377
8.9708 | -92.0382
8.9711 | -92.039
8.9714 | -92.0398
8.9717 | -92.0401
8.9721 | -92.0405
8.9724 | -92.0418
8.9727 | -92.0434
8.973 | -92.0442
8.9733 | -92.0451
8.9736 | -92.0469
8.9739 | -92.0481
8.9742 | -92.0488
8.9745 | -92.0494
8.9749 | -92.0502
8.9752 | -92.0506
8.9755 | -92.051
8.9758 | -92.0516
8.9761 | -92.0522
8.9764 | -92.0529
8.9767 | -92.0538
8.977 | -92.0545
8.9773 | -92.0552
8.9776 | -92.0557
8.978 | -92.0558
8.9783 | -92.0568
8.9786 | -92.0583
8.9789 | -92.0592
8.9792 | -92.06
8.9795 | -92.0606
8.9798 | -92.0602
8.9801 | -92.0606
8.9804 | -92.061
8.9808 | -92.0617
8.9811 | -92.0623
8.9814 | -92.0625
8.9817 | -92.0628
8.982 | -92.0634
8.9823 | -92.0637
8.9826 | -92.0645
8.9829 | -92.0654
8.9832 | -92.0667
8.9835 | -92.0679
8.9839 | -92.0686
8.9842 | -92.0692
8.9845 | -92.0696
8.9848 | -92.0703
8.9851 | -92.0707
8.9854 | -92.0707
8.9857 | -92.0709
8.986 | -92.0714
8.9863 | -92.0724
8.9867 | -92.0732
8.987 | -92.0741
8.9873 | -92.075
8.9876 | -92.075
8.9879 | -92.0751
8.9882 | -92.0754
8.9885 | -92.0762
8.9888 | -92.077
8.9891 | -92.0774
8.9894 | -92.078
8.9898 | -92.0787
8.9901 | -92.0796
8.9904 | -92.0803
8.9907 | -92.0813
8.991 | -92.0821
8.9913 | -92.0825
8.9916 | -92.0831
8.9919 | -92.0836
8.9922 | -92.0837
8.9926 | -92.0842
8.9929 | -92.0849
8.9932 | -92.086
8.9935 | -92.0869
8.9938 | -92.0875
8.9941 | -92.0885
8.9944 | -92.0893
8.9947 | -92.0898
8.995 | -92.0907
8.9953 | -92.0913
8.9957 | -92.0918
8.996 | -92.0923
8.9963 | -92.0929
8.9966 | -92.0936
8.9969 | -92.0942
8.9972 | -92.0947
8.9975 | -92.0954
8.9978 | -92.0957
8.9981 | -92.0966
8.9985 | -92.098
8.9988 | -92.0986
8.9991 | -92.0992
8.9994 | -92.0992
8.9997 | -92.0997
9 | -92.0999
9.0003 | -92.1003
9.0006 | -92.101
9.0009 | -92.1015
9.0012 | -92.102
9.0016 | -92.1026
9.0019 | -92.1031
9.0022 | -92.103
9.0025 | -92.1032
9.0028 | -92.1039
9.0031 | -92.105
9.0034 | -92.1061
9.0037 | -92.1065
9.004 | -92.1069
9.0044 | -92.1058
9.0047 | -92.1057
9.005 | -92.1066
9.0053 | -92.107
9.0056 | -92.1077
9.0059 | -92.1085
9.0062 | -92.1094
9.0065 | -92.11
9.0068 | -92.1107
9.0071 | -92.1114
9.0075 | -92.1119
9.0078 | -92.1123
9.0081 | -92.1129
9.0084 | -92.1138
9.0087 | -92.1145
9.009 | -92.1151
9.0093 | -92.1155
9.0096 | -92.116
9.0099 | -92.1164
9.0103 | -92.1164
9.0106 | -92.1173
9.0109 | -92.1182
9.0112 | -92.1189
9.0115 | -92.1202
9.0118 | -92.1214
9.0121 | -92.1228
9.0124 | -92.1238
9.0127 | -92.1249
9.013 | -92.126
9.0134 | -92.1263
9.0137 | -92.1273
9.014 | -92.1282
9.0143 | -92.129
9.0146 | -92.1294
9.0149 | -92.1299
9.0152 | -92.1305
9.0155 | -92.1311
9.0158 | -92.1318
9.0162 | -92.1325
9.0165 | -92.133
9.0168 | -92.1337
9.0171 | -92.1343
9.0174 | -92.1356
9.0177 | -92.1362
9.018 | -92.1364
9.0183 | -92.137
9.0186 | -92.1381
9.0189 | -92.1393
9.0193 | -92.1399
9.0196 | -92.1403
9.0199 | -92.1409
9.0202 | -92.1417
9.0205 | -92.1422
9.0208 | -92.1429
9.0211 | -92.1438
9.0214 | -92.1447
9.0217 | -92.1446
9.0221 | -92.1458
9.0224 | -92.1463
9.0227 | -92.1476
9.023 | -92.1482
9.0233 | -92.1489
9.0236 | -92.1489
9.0239 | -92.1488
9.0242 | -92.1495
9.0245 | -92.1503
9.0248 | -92.1512
9.0252 | -92.1518
9.0255 | -92.1527
9.0258 | -92.1543
9.0261 | -92.1552
9.0264 | -92.1559
9.0267 | -92.1556
9.027 | -92.1552
9.0273 | -92.1554
9.0276 | -92.1558
9.028 | -92.1562
9.0283 | -92.1568
9.0286 | -92.1571
9.0289 | -92.1574
9.0292 | -92.1578
9.0295 | -92.1583
9.0298 | -92.1585
9.0301 | -92.1587
9.0304 | -92.1592
9.0307 | -92.1597
9.0311 | -92.1601
9.0314 | -92.1602
9.0317 | -92.1608
9.032 | -92.1614
9.0323 | -92.1619
9.0326 | -92.1625
9.0329 | -92.1632
9.0332 | -92.1637
9.0335 | -92.1642
9.0339 | -92.1649
9.0342 | -92.1655
9.0345 | -92.1662
9.0348 | -92.1664
9.0351 | -92.1664
9.0354 | -92.167
9.0357 | -92.1678
9.036 | -92.1683
9.0363 | -92.1683
9.0366 | -92.1687
9.037 | -92.1695
9.0373 | -92.1699
9.0376 | -92.1704
9.0379 | -92.1709
9.0382 | -92.1714
9.0385 | -92.1722
9.0388 | -92.1725
9.0391 | -92.1726
9.0394 | -92.1731
9.0398 | -92.1734
9.0401 | -92.174
9.0404 | -92.1739
9.0407 | -92.1744
9.041 | -92.1748
9.0413 | -92.1753
9.0416 | -92.1755
9.0419 | -92.1761
9.0422 | -92.1765
9.0426 | -92.1769
9.0429 | -92.1773
9.0432 | -92.1776
9.0435 | -92.1785
9.0438 | -92.179
9.0441 | -92.1801
9.0444 | -92.1806
9.0447 | -92.1809
9.045 | -92.1811
9.0453 | -92.1822
9.0457 | -92.183
9.046 | -92.1835
9.0463 | -92.1842
9.0466 | -92.1845
9.0469 | -92.1845
9.0472 | -92.1846
9.0475 | -92.1849
9.0478 | -92.1854
9.0481 | -92.1858
9.0485 | -92.1863
9.0488 | -92.1868
9.0491 | -92.1875
9.0494 | -92.1895
9.0497 | -92.1905
9.05 | -92.191
9.0503 | -92.1917
9.0506 | -92.1923
9.0509 | -92.1928
9.0512 | -92.1932
9.0516 | -92.195
9.0519 | -92.1962
9.0522 | -92.1967
9.0525 | -92.1972
9.0528 | -92.1982
9.0531 | -92.1988
9.0534 | -92.1999
9.0537 | -92.2008
9.054 | -92.2014
9.0543 | -92.2019
9.0547 | -92.2026
9.055 | -92.2034
9.0553 | -92.2035
9.0556 | -92.204
9.0559 | -92.205
9.0562 | -92.2051
9.0565 | -92.2057
9.0568 | -92.2062
9.0571 | -92.2067
9.0575 | -92.2072
9.0578 | -92.2076
9.0581 | -92.2083
9.0584 | -92.2088
9.0587 | -92.2091
9.059 | -92.2099
9.0593 | -92.2109
9.0596 | -92.2117
9.0599 | -92.2122
9.0603 | -92.2128
9.0606 | -92.2138
9.0609 | -92.2149
9.0612 | -92.2164
9.0615 | -92.2175
9.0618 | -92.2178
9.0621 | -92.2181
9.0624 | -92.2176
9.0627 | -92.2178
9.063 | -92.2179
9.0634 | -92.218
9.0637 | -92.2188
9.064 | -92.2193
9.0643 | -92.2199
9.0646 | -92.2204
9.0649 | -92.2206
9.0652 | -92.2219
9.0655 | -92.2228
9.0658 | -92.2234
9.0662 | -92.2236
9.0665 | -92.2238
9.0668 | -92.2242
9.0671 | -92.2247
9.0674 | -92.2259
9.0677 | -92.2269
9.068 | -92.2269
9.0683 | -92.2273
9.0686 | -92.2281
9.0689 | -92.2291
9.0693 | -92.2301
9.0696 | -92.2314
9.0699 | -92.2324
9.0702 | -92.2328
9.0705 | -92.2332
9.0708 | -92.2338
9.0711 | -92.2349
9.0714 | -92.2353
9.0717 | -92.236
9.0721 | -92.2366
9.0724 | -92.2366
9.0727 | -92.2366
9.073 | -92.2371
9.0733 | -92.2382
9.0736 | -92.2388
9.0739 | -92.2392
9.0742 | -92.2395
9.0745 | -92.2396
9.0748 | -92.2409
9.0752 | -92.2423
9.0755 | -92.2426
9.0758 | -92.243
9.0761 | -92.2435
9.0764 | -92.2441
9.0767 | -92.2445
9.077 | -92.2448
9.0773 | -92.2448
9.0776 | -92.2449
9.078 | -92.2453
9.0783 | -92.2456
9.0786 | -92.2463
9.0789 | -92.2472
9.0792 | -92.248
9.0795 | -92.2483
9.0798 | -92.2482
9.0801 | -92.2485
9.0804 | -92.2492
9.0807 | -92.2497
9.0811 | -92.25
9.0814 | -92.2504
9.0817 | -92.2502
9.082 | -92.2502
9.0823 | -92.2507
9.0826 | -92.2511
9.0829 | -92.2521
9.0832 | -92.2531
9.0835 | -92.2533
9.0839 | -92.2538
9.0842 | -92.2541
9.0845 | -92.2543
9.0848 | -92.255
9.0851 | -92.2557
9.0854 | -92.2567
9.0857 | -92.2578
9.086 | -92.2581
9.0863 | -92.2583
9.0866 | -92.2587
9.087 | -92.2587
9.0873 | -92.2592
9.0876 | -92.2596
9.0879 | -92.26
9.0882 | -92.2615
9.0885 | -92.2622
9.0888 | -92.2626
9.0891 | -92.2641
9.0894 | -92.2652
9.0898 | -92.2655
9.0901 | -92.2657
9.0904 | -92.2659
9.0907 | -92.266
9.091 | -92.266
9.0913 | -92.2666
9.0916 | -92.2674
9.0919 | -92.2687
9.0922 | -92.2693
9.0925 | -92.2693
9.0929 | -92.2692
9.0932 | -92.2695
9.0935 | -92.2703
9.0938 | -92.2705
9.0941 | -92.271
9.0944 | -92.272
9.0947 | -92.2729
9.095 | -92.273
9.0953 | -92.2729
9.0957 | -92.2734
9.096 | -92.274
9.0963 | -92.275
9.0966 | -92.2756
9.0969 | -92.2759
9.0972 | -92.276
9.0975 | -92.2762
9.0978 | -92.2765
9.0981 | -92.277
9.0984 | -92.2774
9.0988 | -92.2779
9.0991 | -92.2783
9.0994 | -92.2784
9.0997 | -92.2784
9.1 | -92.2784
9.1003 | -92.2789
9.1006 | -92.2799
9.1009 | -92.2803
9.1012 | -92.2806
9.1016 | -92.281
9.1019 | -92.2815
9.1022 | -92.2821
9.1025 | -92.2827
9.1028 | -92.283
9.1031 | -92.2836
9.1034 | -92.2837
9.1037 | -92.2843
9.104 | -92.2851
9.1043 | -92.2851
9.1047 | -92.2855
9.105 | -92.2862
9.1053 | -92.2866
9.1056 | -92.2868
9.1059 | -92.2869
9.1062 | -92.2873
9.1065 | -92.2879
9.1068 | -92.2885
9.1071 | -92.2886
9.1075 | -92.2888
9.1078 | -92.2889
9.1081 | -92.289
9.1084 | -92.2894
9.1087 | -92.2898
9.109 | -92.2899
9.1093 | -92.2901
9.1096 | -92.2906
9.1099 | -92.2907
9.1102 | -92.2906
9.1106 | -92.2909
9.1109 | -92.2913
9.1112 | -92.2918
9.1115 | -92.2915
9.1118 | -92.2907
9.1121 | -92.2907
9.1124 | -92.291
9.1127 | -92.2914
9.113 | -92.2915
9.1134 | -92.2917
9.1137 | -92.2919
9.114 | -92.2918
9.1143 | -92.2922
9.1146 | -92.2924
9.1149 | -92.2927
9.1152 | -92.293
9.1155 | -92.2934
9.1158 | -92.2938
9.1161 | -92.2942
9.1165 | -92.2945
9.1168 | -92.295
9.1171 | -92.2952
9.1174 | -92.2954
9.1177 | -92.2956
9.118 | -92.2959
9.1183 | -92.2962
9.1186 | -92.2965
9.1189 | -92.2969
9.1193 | -92.2969
9.1196 | -92.297
9.1199 | -92.2974
9.1202 | -92.2974
9.1205 | -92.2976
9.1208 | -92.2977
9.1211 | -92.2975
9.1214 | -92.2975
9.1217 | -92.2975
9.122 | -92.2977
9.1224 | -92.2985
9.1227 | -92.2994
9.123 | -92.2997
9.1233 | -92.2997
9.1236 | -92.2996
9.1239 | -92.2997
9.1242 | -92.2999
9.1245 | -92.3003
9.1248 | -92.3007
9.1252 | -92.3007
9.1255 | -92.3009
9.1258 | -92.3012
9.1261 | -92.3009
9.1264 | -92.3009
9.1267 | -92.3014
9.127 | -92.3019
9.1273 | -92.302
9.1276 | -92.3022
9.1279 | -92.3025
9.1283 | -92.3029
9.1286 | -92.3031
9.1289 | -92.3034
9.1292 | -92.3036
9.1295 | -92.3037
9.1298 | -92.3039
9.1301 | -92.3042
9.1304 | -92.3043
9.1307 | -92.3049
9.1311 | -92.3062
9.1314 | -92.3074
9.1317 | -92.3078
9.132 | -92.3086
9.1323 | -92.3092
9.1326 | -92.3092
9.1329 | -92.3099
9.1332 | -92.3111
9.1335 | -92.3115
9.1338 | -92.3111
9.1342 | -92.311
9.1345 | -92.3113
9.1348 | -92.3115
9.1351 | -92.3114
9.1354 | -92.3117
9.1357 | -92.312
9.136 | -92.3122
9.1363 | -92.3123
9.1366 | -92.3126
9.137 | -92.3132
9.1373 | -92.313
9.1376 | -92.313
9.1379 | -92.3138
9.1382 | -92.3141
9.1385 | -92.3144
9.1388 | -92.3149
9.1391 | -92.3154
9.1394 | -92.3163
9.1397 | -92.3174
9.1401 | -92.3186
9.1404 | -92.3188
9.1407 | -92.3188
9.141 | -92.3191
9.1413 | -92.3195
9.1416 | -92.3197
9.1419 | -92.32
9.1422 | -92.3204
9.1425 | -92.3209
9.1429 | -92.3214
9.1432 | -92.3217
9.1435 | -92.3218
9.1438 | -92.3221
9.1441 | -92.3233
9.1444 | -92.324
9.1447 | -92.324
9.145 | -92.3235
9.1453 | -92.3234
9.1456 | -92.3236
9.146 | -92.3238
9.1463 | -92.3242
9.1466 | -92.3245
9.1469 | -92.3248
9.1472 | -92.325
9.1475 | -92.3251
9.1478 | -92.3255
9.1481 | -92.3256
9.1484 | -92.3261
9.1488 | -92.3267
9.1491 | -92.3271
9.1494 | -92.3276
9.1497 | -92.3282
9.15 | -92.3287
9.1503 | -92.3294
9.1506 | -92.3297
9.1509 | -92.3298
9.1512 | -92.3296
9.1515 | -92.3297
9.1519 | -92.3303
9.1522 | -92.33
9.1525 | -92.3298
9.1528 | -92.3301
9.1531 | -92.3306
9.1534 | -92.3314
9.1537 | -92.3317
9.154 | -92.332
9.1543 | -92.332
9.1547 | -92.3323
9.155 | -92.3326
9.1553 | -92.3329
9.1556 | -92.3327
9.1559 | -92.3329
9.1562 | -92.3328
9.1565 | -92.3329
9.1568 | -92.3335
9.1571 | -92.3335
9.1574 | -92.3335
9.1578 | -92.3335
9.1581 | -92.3337
9.1584 | -92.334
9.1587 | -92.3344
9.159 | -92.3345
9.1593 | -92.3348
9.1596 | -92.3357
9.1599 | -92.3363
9.1602 | -92.3362
9.1606 | -92.337
9.1609 | -92.338
9.1612 | -92.3385
9.1615 | -92.3392
9.1618 | -92.3401
9.1621 | -92.3406
9.1624 | -92.3408
9.1627 | -92.3419
9.163 | -92.3426
9.1633 | -92.3425
9.1637 | -92.3426
9.164 | -92.343
9.1643 | -92.3433
9.1646 | -92.3439
9.1649 | -92.3442
9.1652 | -92.3442
9.1655 | -92.3445
9.1658 | -92.3451
9.1661 | -92.3453
9.1665 | -92.3456
9.1668 | -92.3457
9.1671 | -92.3458
9.1674 | -92.3465
9.1677 | -92.3471
9.168 | -92.3473
9.1683 | -92.3475
9.1686 | -92.3479
9.1689 | -92.3483
9.1692 | -92.3486
9.1696 | -92.3493
9.1699 | -92.3499
9.1702 | -92.3504
9.1705 | -92.3503
9.1708 | -92.3507
9.1711 | -92.3513
9.1714 | -92.3523
9.1717 | -92.3525
9.172 | -92.3528
9.1724 | -92.3526
9.1727 | -92.3529
9.173 | -92.3532
9.1733 | -92.3534
9.1736 | -92.3539
9.1739 | -92.3544
9.1742 | -92.3544
9.1745 | -92.3542
9.1748 | -92.3544
9.1752 | -92.3548
9.1755 | -92.3548
9.1758 | -92.3546
9.1761 | -92.3548
9.1764 | -92.3543
9.1767 | -92.354
9.177 | -92.3543
9.1773 | -92.3542
9.1776 | -92.3539
9.1779 | -92.3543
9.1783 | -92.3547
9.1786 | -92.3551
9.1789 | -92.3554
9.1792 | -92.3556
9.1795 | -92.3562
9.1798 | -92.3574
9.1801 | -92.3581
9.1804 | -92.3581
9.1807 | -92.3581
9.1811 | -92.3585
9.1814 | -92.3596
9.1817 | -92.3611
9.182 | -92.3613
9.1823 | -92.3617
9.1826 | -92.3619
9.1829 | -92.3624
9.1832 | -92.3631
9.1835 | -92.3635
9.1838 | -92.3633
9.1842 | -92.3636
9.1845 | -92.3641
9.1848 | -92.3645
9.1851 | -92.3648
9.1854 | -92.3639
9.1857 | -92.3634
9.186 | -92.3636
9.1863 | -92.3634
9.1866 | -92.3637
9.1869 | -92.3645
9.1873 | -92.3651
9.1876 | -92.3653
9.1879 | -92.3656
9.1882 | -92.3659
9.1885 | -92.3661
9.1888 | -92.366
9.1891 | -92.3663
9.1894 | -92.3669
9.1897 | -92.3673
9.1901 | -92.3675
9.1904 | -92.3677
9.1907 | -92.3682
9.191 | -92.3685
9.1913 | -92.3687
9.1916 | -92.3689
9.1919 | -92.3692
9.1922 | -92.3692
9.1925 | -92.3697
9.1928 | -92.3697
9.1932 | -92.3699
9.1935 | -92.3699
9.1938 | -92.3701
9.1941 | -92.3705
9.1944 | -92.3703
9.1947 | -92.3702
9.195 | -92.3704
9.1953 | -92.3705
9.1956 | -92.3704
9.196 | -92.3706
9.1963 | -92.3708
9.1966 | -92.3716
9.1969 | -92.3728
9.1972 | -92.3735
9.1975 | -92.3738
9.1978 | -92.3743
9.1981 | -92.3747
9.1984 | -92.3751
9.1988 | -92.3755
9.1991 | -92.3758
9.1994 | -92.376
9.1997 | -92.3763
9.2 | -92.3766
9.2003 | -92.3771
9.2006 | -92.3779
9.2009 | -92.3783
9.2012 | -92.3786
9.2015 | -92.3791
9.2019 | -92.3795
9.2022 | -92.3797
9.2025 | -92.38
9.2028 | -92.3806
9.2031 | -92.3815
9.2034 | -92.382
9.2037 | -92.3823
9.204 | -92.3827
9.2043 | -92.3829
9.2047 | -92.383
9.205 | -92.3833
9.2053 | -92.3838
9.2056 | -92.3842
9.2059 | -92.3847
9.2062 | -92.385
9.2065 | -92.3852
9.2068 | -92.3852
9.2071 | -92.3849
9.2074 | -92.385
9.2078 | -92.3851
9.2081 | -92.3854
9.2084 | -92.3862
9.2087 | -92.3865
9.209 | -92.3868
9.2093 | -92.3872
9.2096 | -92.3873
9.2099 | -92.3877
9.2102 | -92.3882
9.2106 | -92.3876
9.2109 | -92.388
9.2112 | -92.389
9.2115 | -92.3899
9.2118 | -92.3907
9.2121 | -92.3914
9.2124 | -92.3922
9.2127 | -92.3938
9.213 | -92.3958
9.2133 | -92.3973
9.2137 | -92.3984
9.214 | -92.3993
9.2143 | -92.4
9.2146 | -92.4003
9.2149 | -92.4001
9.2152 | -92.4
9.2155 | -92.4002
9.2158 | -92.4004
9.2161 | -92.4007
9.2165 | -92.4012
9.2168 | -92.4018
9.2171 | -92.4032
9.2174 | -92.4044
9.2177 | -92.4049
9.218 | -92.4049
9.2183 | -92.405
9.2186 | -92.4057
9.2189 | -92.4057
9.2192 | -92.4062
9.2196 | -92.4068
9.2199 | -92.4076
9.2202 | -92.4077
9.2205 | -92.4076
9.2208 | -92.4084
9.2211 | -92.4093
9.2214 | -92.4101
9.2217 | -92.4107
9.222 | -92.4112
9.2224 | -92.4115
9.2227 | -92.4119
9.223 | -92.4121
9.2233 | -92.4121
9.2236 | -92.4125
9.2239 | -92.4136
9.2242 | -92.4145
9.2245 | -92.4149
9.2248 | -92.4153
9.2251 | -92.4161
9.2255 | -92.4168
9.2258 | -92.4171
9.2261 | -92.4178
9.2264 | -92.4185
9.2267 | -92.4187
9.227 | -92.4188
9.2273 | -92.419
9.2276 | -92.4194
9.2279 | -92.4202
9.2283 | -92.4214
9.2286 | -92.4219
9.2289 | -92.4221
9.2292 | -92.4217
9.2295 | -92.4216
9.2298 | -92.4217
9.2301 | -92.4223
9.2304 | -92.4232
9.2307 | -92.4237
9.231 | -92.4243
9.2314 | -92.4246
9.2317 | -92.4252
9.232 | -92.4261
9.2323 | -92.4266
9.2326 | -92.4269
9.2329 | -92.4269
9.2332 | -92.4275
9.2335 | -92.4283
9.2338 | -92.4285
9.2342 | -92.4284
9.2345 | -92.4285
9.2348 | -92.429
9.2351 | -92.4302
9.2354 | -92.4309
9.2357 | -92.431
9.236 | -92.4311
9.2363 | -92.4315
9.2366 | -92.432
9.2369 | -92.432
9.2373 | -92.4328
9.2376 | -92.4331
9.2379 | -92.4334
9.2382 | -92.434
9.2385 | -92.4343
9.2388 | -92.4342
9.2391 | -92.4343
9.2394 | -92.4346
9.2397 | -92.4352
9.2401 | -92.4352
9.2404 | -92.4357
9.2407 | -92.4367
9.241 | -92.4367
9.2413 | -92.4371
9.2416 | -92.4384
9.2419 | -92.4397
9.2422 | -92.4398
9.2425 | -92.4399
9.2428 | -92.4402
9.2432 | -92.4406
9.2435 | -92.441
9.2438 | -92.4414
9.2441 | -92.4417
9.2444 | -92.4416
9.2447 | -92.4418
9.245 | -92.4418
9.2453 | -92.4418
9.2456 | -92.4419
9.246 | -92.4423
9.2463 | -92.443
9.2466 | -92.4432
9.2469 | -92.4432
9.2472 | -92.4432
9.2475 | -92.4436
9.2478 | -92.4443
9.2481 | -92.4451
9.2484 | -92.4456
9.2487 | -92.4465
9.2491 | -92.4471
9.2494 | -92.4475
9.2497 | -92.4477
9.25 | -92.4481
9.2503 | -92.4481
9.2506 | -92.4485
9.2509 | -92.449
9.2512 | -92.4493
9.2515 | -92.4494
9.2519 | -92.4496
9.2522 | -92.4496
9.2525 | -92.449
9.2528 | -92.4492
9.2531 | -92.45
9.2534 | -92.451
9.2537 | -92.4518
9.254 | -92.4525
9.2543 | -92.4534
9.2546 | -92.4545
9.255 | -92.455
9.2553 | -92.4559
9.2556 | -92.4563
9.2559 | -92.4564
9.2562 | -92.4571
9.2565 | -92.4585
9.2568 | -92.4589
9.2571 | -92.4591
9.2574 | -92.4589
9.2578 | -92.4589
9.2581 | -92.4592
9.2584 | -92.4595
9.2587 | -92.4597
9.259 | -92.461
9.2593 | -92.4619
9.2596 | -92.4623
9.2599 | -92.4627
9.2602 | -92.4637
9.2605 | -92.4641
9.2609 | -92.4645
9.2612 | -92.4651
9.2615 | -92.4655
9.2618 | -92.4659
9.2621 | -92.466
9.2624 | -92.4662
9.2627 | -92.4666
9.263 | -92.4669
9.2633 | -92.4675
9.2637 | -92.4675
9.264 | -92.4679
9.2643 | -92.4686
9.2646 | -92.4689
9.2649 | -92.4695
9.2652 | -92.4697
9.2655 | -92.4698
9.2658 | -92.4716
9.2661 | -92.4732
9.2664 | -92.4737
9.2668 | -92.4739
9.2671 | -92.4739
9.2674 | -92.4743
9.2677 | -92.475
9.268 | -92.476
9.2683 | -92.477
9.2686 | -92.4776
9.2689 | -92.4781
9.2692 | -92.4783
9.2696 | -92.4786
9.2699 | -92.4789
9.2702 | -92.4791
9.2705 | -92.4788
9.2708 | -92.4795
9.2711 | -92.48
9.2714 | -92.48
9.2717 | -92.4808
9.272 | -92.4813
9.2723 | -92.4819
9.2727 | -92.4828
9.273 | -92.4838
9.2733 | -92.4844
9.2736 | -92.4845
9.2739 | -92.4849
9.2742 | -92.4853
9.2745 | -92.4861
9.2748 | -92.487
9.2751 | -92.4874
9.2755 | -92.4878
9.2758 | -92.4882
9.2761 | -92.4887
9.2764 | -92.4892
9.2767 | -92.4897
9.277 | -92.4902
9.2773 | -92.4915
9.2776 | -92.4921
9.2779 | -92.4922
9.2782 | -92.4928
9.2786 | -92.4934
9.2789 | -92.4941
9.2792 | -92.4947
9.2795 | -92.4952
9.2798 | -92.496
9.2801 | -92.4964
9.2804 | -92.4968
9.2807 | -92.497
9.281 | -92.4974
9.2814 | -92.4959
9.2817 | -92.4923
9.282 | -92.4922
9.2823 | -92.4926
9.2826 | -92.493
9.2829 | -92.4934
9.2832 | -92.4936
9.2835 | -92.4938
9.2838 | -92.4944
9.2841 | -92.4947
9.2845 | -92.4949
9.2848 | -92.4953
9.2851 | -92.4953
9.2854 | -92.4957
9.2857 | -92.4965
9.286 | -92.4975
9.2863 | -92.498
9.2866 | -92.4984
9.2869 | -92.4987
9.2873 | -92.4983
9.2876 | -92.4979
9.2879 | -92.4986
9.2882 | -92.4994
9.2885 | -92.5001
9.2888 | -92.5005
9.2891 | -92.5009
9.2894 | -92.5015
9.2897 | -92.5023
9.29 | -92.5033
9.2904 | -92.5041
9.2907 | -92.5044
9.291 | -92.5054
9.2913 | -92.5059
9.2916 | -92.5066
9.2919 | -92.5075
9.2922 | -92.5084
9.2925 | -92.509
9.2928 | -92.5098
9.2932 | -92.5104
9.2935 | -92.5103
9.2938 | -92.5106
9.2941 | -92.5113
9.2944 | -92.5121
9.2947 | -92.5128
9.295 | -92.514
9.2953 | -92.5148
9.2956 | -92.5153
9.2959 | -92.5156
9.2963 | -92.5161
9.2966 | -92.5164
9.2969 | -92.5172
9.2972 | -92.5176
9.2975 | -92.5181
9.2978 | -92.5184
9.2981 | -92.5186
9.2984 | -92.5189
9.2987 | -92.5195
9.2991 | -92.5201
9.2994 | -92.5207
9.2997 | -92.5214
9.3 | -92.5221
9.3003 | -92.5228
9.3006 | -92.5233
9.3009 | -92.5235
9.3012 | -92.524
9.3015 | -92.5246
9.3018 | -92.5253
9.3022 | -92.5263
9.3025 | -92.5271
9.3028 | -92.5275
9.3031 | -92.5276
9.3034 | -92.5273
9.3037 | -92.5274
9.304 | -92.5278
9.3043 | -92.5288
9.3046 | -92.5301
9.305 | -92.5308
9.3053 | -92.5314
9.3056 | -92.532
9.3059 | -92.5321
9.3062 | -92.5322
9.3065 | -92.5328
9.3068 | -92.5333
9.3071 | -92.5338
9.3074 | -92.5343
9.3077 | -92.5346
9.3081 | -92.5358
9.3084 | -92.5376
9.3087 | -92.5381
9.309 | -92.5391
9.3093 | -92.5394
9.3096 | -92.5399
9.3099 | -92.5407
9.3102 | -92.541
9.3105 | -92.5414
9.3109 | -92.5424
9.3112 | -92.5427
9.3115 | -92.5432
9.3118 | -92.5431
9.3121 | -92.5433
9.3124 | -92.5437
9.3127 | -92.5438
9.313 | -92.5442
9.3133 | -92.5446
9.3136 | -92.5452
9.314 | -92.5457
9.3143 | -92.5463
9.3146 | -92.547
9.3149 | -92.5476
9.3152 | -92.5485
9.3155 | -92.549
9.3158 | -92.5496
9.3161 | -92.5502
9.3164 | -92.5508
9.3168 | -92.5512
9.3171 | -92.5518
9.3174 | -92.5525
9.3177 | -92.5529
9.318 | -92.5539
9.3183 | -92.5545
9.3186 | -92.5547
9.3189 | -92.5552
9.3192 | -92.5556
9.3195 | -92.5561
9.3199 | -92.5567
9.3202 | -92.5571
9.3205 | -92.5576
9.3208 | -92.5582
9.3211 | -92.559
9.3214 | -92.5598
9.3217 | -92.5604
9.322 | -92.5612
9.3223 | -92.5622
9.3227 | -92.5633
9.323 | -92.5641
9.3233 | -92.5647
9.3236 | -92.5653
9.3239 | -92.5667
9.3242 | -92.5685
9.3245 | -92.5699
9.3248 | -92.5708
9.3251 | -92.5714
9.3255 | -92.5718
9.3258 | -92.5728
9.3261 | -92.5735
9.3264 | -92.5731
9.3267 | -92.5731
9.327 | -92.5741
9.3273 | -92.5748
9.3276 | -92.5752
9.3279 | -92.5758
9.3282 | -92.5766
9.3286 | -92.5772
9.3289 | -92.5778
9.3292 | -92.5785
9.3295 | -92.5797
9.3298 | -92.5809
9.3301 | -92.5816
9.3304 | -92.5824
9.3307 | -92.5832
9.331 | -92.5838
9.3314 | -92.5844
9.3317 | -92.5851
9.332 | -92.5858
9.3323 | -92.5864
9.3326 | -92.5873
9.3329 | -92.5878
9.3332 | -92.5885
9.3335 | -92.5892
9.3338 | -92.5899
9.3341 | -92.5907
9.3345 | -92.5918
9.3348 | -92.5931
9.3351 | -92.5938
9.3354 | -92.5941
9.3357 | -92.5945
9.336 | -92.5952
9.3363 | -92.596
9.3366 | -92.5967
9.3369 | -92.5973
9.3372 | -92.598
9.3376 | -92.5989
9.3379 | -92.5989
9.3382 | -92.5987
9.3385 | -92.5995
9.3388 | -92.6004
9.3391 | -92.6009
9.3394 | -92.6013
9.3397 | -92.6016
9.34 | -92.6022
9.3404 | -92.6028
9.3407 | -92.6035
9.341 | -92.6038
9.3413 | -92.6042
9.3416 | -92.6048
9.3419 | -92.6055
9.3422 | -92.6063
9.3425 | -92.6071
9.3428 | -92.608
9.3431 | -92.6089
9.3435 | -92.6093
9.3438 | -92.6105
9.3441 | -92.6118
9.3444 | -92.6123
9.3447 | -92.6128
9.345 | -92.6136
9.3453 | -92.614
9.3456 | -92.6146
9.3459 | -92.6153
9.3463 | -92.6158
9.3466 | -92.6161
9.3469 | -92.6163
9.3472 | -92.6168
9.3475 | -92.6175
9.3478 | -92.6181
9.3481 | -92.6188
9.3484 | -92.6195
9.3487 | -92.62
9.3491 | -92.6207
9.3494 | -92.6218
9.3497 | -92.6227
9.35 | -92.6234
9.3503 | -92.6239
9.3506 | -92.6247
9.3509 | -92.6254
9.3512 | -92.6267
9.3515 | -92.6286
9.3518 | -92.6299
9.3522 | -92.6301
9.3525 | -92.6308
9.3528 | -92.6314
9.3531 | -92.6319
9.3534 | -92.6328
9.3537 | -92.6339
9.354 | -92.6357
9.3543 | -92.6374
9.3546 | -92.6379
9.355 | -92.6393
9.3553 | -92.6402
9.3556 | -92.6404
9.3559 | -92.6409
9.3562 | -92.6415
9.3565 | -92.6423
9.3568 | -92.643
9.3571 | -92.6434
9.3574 | -92.6431
9.3577 | -92.6429
9.3581 | -92.6436
9.3584 | -92.6449
9.3587 | -92.6456
9.359 | -92.6462
9.3593 | -92.6469
9.3596 | -92.6476
9.3599 | -92.648
9.3602 | -92.6487
9.3605 | -92.6492
9.3609 | -92.6502
9.3612 | -92.6511
9.3615 | -92.6512
9.3618 | -92.6519
9.3621 | -92.6525
9.3624 | -92.6533
9.3627 | -92.6535
9.363 | -92.6541
9.3633 | -92.6545
9.3636 | -92.6542
9.364 | -92.6545
9.3643 | -92.6551
9.3646 | -92.6559
9.3649 | -92.6569
9.3652 | -92.6577
9.3655 | -92.6582
9.3658 | -92.6588
9.3661 | -92.6591
9.3664 | -92.6595
9.3668 | -92.6605
9.3671 | -92.6613
9.3674 | -92.6621
9.3677 | -92.663
9.368 | -92.6635
9.3683 | -92.6632
9.3686 | -92.6639
9.3689 | -92.6646
9.3692 | -92.6654
9.3695 | -92.6668
9.3699 | -92.6686
9.3702 | -92.6696
9.3705 | -92.6703
9.3708 | -92.671
9.3711 | -92.6713
9.3714 | -92.6717
9.3717 | -92.6725
9.372 | -92.6731
9.3723 | -92.6734
9.3727 | -92.674
9.373 | -92.6746
9.3733 | -92.6752
9.3736 | -92.6758
9.3739 | -92.6762
9.3742 | -92.677
9.3745 | -92.6781
9.3748 | -92.6793
9.3751 | -92.6801
9.3754 | -92.6809
9.3758 | -92.682
9.3761 | -92.6826
9.3764 | -92.683
9.3767 | -92.6841
9.377 | -92.6853
9.3773 | -92.6859
9.3776 | -92.6871
9.3779 | -92.6884
9.3782 | -92.6893
9.3786 | -92.6901
9.3789 | -92.6907
9.3792 | -92.6912
9.3795 | -92.692
9.3798 | -92.6927
9.3801 | -92.6931
9.3804 | -92.6936
9.3807 | -92.6938
9.381 | -92.6946
9.3813 | -92.6954
9.3817 | -92.6961
9.382 | -92.697
9.3823 | -92.6977
9.3826 | -92.6985
9.3829 | -92.6995
9.3832 | -92.7005
9.3835 | -92.7012
9.3838 | -92.7017
9.3841 | -92.7016
9.3845 | -92.7012
9.3848 | -92.7015
9.3851 | -92.7021
9.3854 | -92.7025
9.3857 | -92.7035
9.386 | -92.7049
9.3863 | -92.7058
9.3866 | -92.7063
9.3869 | -92.7072
9.3872 | -92.7087
9.3876 | -92.7091
9.3879 | -92.7099
9.3882 | -92.7107
9.3885 | -92.7118
9.3888 | -92.7124
9.3891 | -92.7132
9.3894 | -92.7134
9.3897 | -92.7136
9.39 | -92.7141
9.3904 | -92.7147
9.3907 | -92.7152
9.391 | -92.7164
9.3913 | -92.7179
9.3916 | -92.7196
9.3919 | -92.7218
9.3922 | -92.7225
9.3925 | -92.7236
9.3928 | -92.7245
9.3931 | -92.7252
9.3935 | -92.726
9.3938 | -92.7266
9.3941 | -92.7272
9.3944 | -92.7281
9.3947 | -92.7286
9.395 | -92.7292
9.3953 | -92.73
9.3956 | -92.7304
9.3959 | -92.7309
9.3963 | -92.7316
9.3966 | -92.7325
9.3969 | -92.7333
9.3972 | -92.7339
9.3975 | -92.7345
9.3978 | -92.735
9.3981 | -92.7357
9.3984 | -92.7365
9.3987 | -92.7371
9.399 | -92.7375
9.3994 | -92.7379
9.3997 | -92.7384
9.4 | -92.7391
9.4003 | -92.7398
9.4006 | -92.7404
9.4009 | -92.7411
9.4012 | -92.7419
9.4015 | -92.7425
9.4018 | -92.7431
9.4022 | -92.7436
9.4025 | -92.7447
9.4028 | -92.7455
9.4031 | -92.7459
9.4034 | -92.7464
9.4037 | -92.7475
9.404 | -92.7481
9.4043 | -92.7487
9.4046 | -92.7497
9.4049 | -92.7501
9.4053 | -92.7517
9.4056 | -92.7529
9.4059 | -92.7535
9.4062 | -92.7542
9.4065 | -92.755
9.4068 | -92.7555
9.4071 | -92.7562
9.4074 | -92.7573
9.4077 | -92.7579
9.4081 | -92.7588
9.4084 | -92.7597
9.4087 | -92.7612
9.409 | -92.762
9.4093 | -92.7627
9.4096 | -92.7633
9.4099 | -92.7638
9.4102 | -92.7642
9.4105 | -92.7647
9.4108 | -92.7654
9.4112 | -92.7662
9.4115 | -92.7675
9.4118 | -92.7686
9.4121 | -92.7696
9.4124 | -92.7703
9.4127 | -92.7715
9.413 | -92.7728
9.4133 | -92.7736
9.4136 | -92.7741
9.414 | -92.7744
9.4143 | -92.7752
9.4146 | -92.7763
9.4149 | -92.7769
9.4152 | -92.7775
9.4155 | -92.7781
9.4158 | -92.7787
9.4161 | -92.7794
9.4164 | -92.7798
9.4167 | -92.7805
9.4171 | -92.7813
9.4174 | -92.7817
9.4177 | -92.7821
9.418 | -92.783
9.4183 | -92.7842
9.4186 | -92.785
9.4189 | -92.7859
9.4192 | -92.7867
9.4195 | -92.7874
9.4199 | -92.7882
9.4202 | -92.7887
9.4205 | -92.7891
9.4208 | -92.7896
9.4211 | -92.7901
9.4214 | -92.7905
9.4217 | -92.7912
9.422 | -92.7915
9.4223 | -92.7918
9.4226 | -92.7924
9.423 | -92.7932
9.4233 | -92.7942
9.4236 | -92.7951
9.4239 | -92.7958
9.4242 | -92.7961
9.4245 | -92.7968
9.4248 | -92.7975
9.4251 | -92.7983
9.4254 | -92.7989
9.4258 | -92.7995
9.4261 | -92.8002
9.4264 | -92.801
9.4267 | -92.8017
9.427 | -92.8027
9.4273 | -92.8036
9.4276 | -92.8043
9.4279 | -92.8049
9.4282 | -92.8055
9.4285 | -92.8064
9.4289 | -92.8065
9.4292 | -92.8072
9.4295 | -92.8083
9.4298 | -92.8091
9.4301 | -92.8099
9.4304 | -92.8109
9.4307 | -92.8114
9.431 | -92.8118
9.4313 | -92.8125
9.4317 | -92.813
9.432 | -92.8135
9.4323 | -92.8141
9.4326 | -92.8147
9.4329 | -92.8149
9.4332 | -92.8154
9.4335 | -92.8161
9.4338 | -92.8172
9.4341 | -92.8177
9.4344 | -92.818
9.4348 | -92.8187
9.4351 | -92.8189
9.4354 | -92.8205
9.4357 | -92.8221
9.436 | -92.8233
9.4363 | -92.8245
9.4366 | -92.8252
9.4369 | -92.8261
9.4372 | -92.8275
9.4376 | -92.8289
9.4379 | -92.8302
9.4382 | -92.8312
9.4385 | -92.8322
9.4388 | -92.8334
9.4391 | -92.8342
9.4394 | -92.8349
9.4397 | -92.8354
9.44 | -92.8363
9.4403 | -92.8379
9.4407 | -92.8391
9.441 | -92.8403
9.4413 | -92.8423
9.4416 | -92.8432
9.4419 | -92.8436
9.4422 | -92.8443
9.4425 | -92.8445
9.4428 | -92.8448
9.4431 | -92.8455
9.4435 | -92.8464
9.4438 | -92.8473
9.4441 | -92.8479
9.4444 | -92.8485
9.4447 | -92.849
9.445 | -92.8494
9.4453 | -92.8496
9.4456 | -92.8497
9.4459 | -92.85
9.4462 | -92.8509
9.4466 | -92.8522
9.4469 | -92.8532
9.4472 | -92.854
9.4475 | -92.8545
9.4478 | -92.8555
9.4481 | -92.8559
9.4484 | -92.8565
9.4487 | -92.8571
9.449 | -92.8568
9.4494 | -92.8572
9.4497 | -92.8576
9.45 | -92.8581
9.4503 | -92.8584
9.4506 | -92.8588
9.4509 | -92.8593
9.4512 | -92.8602
9.4515 | -92.861
9.4518 | -92.8614
9.4521 | -92.8609
9.4525 | -92.8601
9.4528 | -92.8603
9.4531 | -92.8604
9.4534 | -92.8609
9.4537 | -92.8614
9.454 | -92.8624
9.4543 | -92.863
9.4546 | -92.8636
9.4549 | -92.8642
9.4553 | -92.865
9.4556 | -92.8655
9.4559 | -92.866
9.4562 | -92.8666
9.4565 | -92.867
9.4568 | -92.8676
9.4571 | -92.8684
9.4574 | -92.8694
9.4577 | -92.8707
9.458 | -92.8712
9.4584 | -92.8721
9.4587 | -92.8727
9.459 | -92.8733
9.4593 | -92.8743
9.4596 | -92.8752
9.4599 | -92.8761
9.4602 | -92.8768
9.4605 | -92.8773
9.4608 | -92.8778
9.4612 | -92.8784
9.4615 | -92.8792
9.4618 | -92.8798
9.4621 | -92.8804
9.4624 | -92.8808
9.4627 | -92.8814
9.463 | -92.882
9.4633 | -92.8826
9.4636 | -92.8831
9.464 | -92.8842
9.4643 | -92.8848
9.4646 | -92.8853
9.4649 | -92.8862
9.4652 | -92.8868
9.4655 | -92.8871
9.4658 | -92.8873
9.4661 | -92.8877
9.4664 | -92.8882
9.4667 | -92.8888
9.4671 | -92.8894
9.4674 | -92.8897
9.4677 | -92.8901
9.468 | -92.8905
9.4683 | -92.8911
9.4686 | -92.891
9.4689 | -92.891
9.4692 | -92.8914
9.4695 | -92.8915
9.4698 | -92.8913
9.4702 | -92.8916
9.4705 | -92.8921
9.4708 | -92.8924
9.4711 | -92.8932
9.4714 | -92.8937
9.4717 | -92.8941
9.472 | -92.8943
9.4723 | -92.8946
9.4726 | -92.8951
9.473 | -92.8957
9.4733 | -92.8965
9.4736 | -92.8973
9.4739 | -92.898
9.4742 | -92.8987
9.4745 | -92.8991
9.4748 | -92.8994
9.4751 | -92.8997
9.4754 | -92.9001
9.4757 | -92.9006
9.4761 | -92.9011
9.4764 | -92.9019
9.4767 | -92.903
9.477 | -92.9036
9.4773 | -92.9044
9.4776 | -92.9047
9.4779 | -92.9055
9.4782 | -92.9065
9.4785 | -92.9077
9.4789 | -92.9093
9.4792 | -92.9095
9.4795 | -92.9105
9.4798 | -92.911
9.4801 | -92.9114
9.4804 | -92.9117
9.4807 | -92.9116
9.481 | -92.912
9.4813 | -92.9126
9.4817 | -92.9132
9.482 | -92.9138
9.4823 | -92.9144
9.4826 | -92.9148
9.4829 | -92.9156
9.4832 | -92.9165
9.4835 | -92.9175
9.4838 | -92.9181
9.4841 | -92.9187
9.4844 | -92.9193
9.4848 | -92.9199
9.4851 | -92.9205
9.4854 | -92.921
9.4857 | -92.9215
9.486 | -92.922
9.4863 | -92.9225
9.4866 | -92.9237
9.4869 | -92.9248
9.4872 | -92.9252
9.4876 | -92.9256
9.4879 | -92.9263
9.4882 | -92.9272
9.4885 | -92.9282
9.4888 | -92.9288
9.4891 | -92.9292
9.4894 | -92.9297
9.4897 | -92.9303
9.49 | -92.9318
9.4903 | -92.933
9.4907 | -92.9335
9.491 | -92.934
9.4913 | -92.9343
9.4916 | -92.9344
9.4919 | -92.9348
9.4922 | -92.9353
9.4925 | -92.9358
9.4928 | -92.9366
9.4931 | -92.9371
9.4935 | -92.9377
9.4938 | -92.9384
9.4941 | -92.9389
9.4944 | -92.9395
9.4947 | -92.9401
9.495 | -92.9401
9.4953 | -92.9401
9.4956 | -92.9407
9.4959 | -92.9412
9.4962 | -92.9418
9.4966 | -92.9426
9.4969 | -92.9429
9.4972 | -92.9434
9.4975 | -92.9441
9.4978 | -92.9452
9.4981 | -92.9457
9.4984 | -92.9461
9.4987 | -92.9467
9.499 | -92.9469
9.4994 | -92.9469
9.4997 | -92.9471
9.5 | -92.9474
9.5003 | -92.9479
9.5006 | -92.948
9.5009 | -92.9475
9.5012 | -92.9477
9.5015 | -92.9479
9.5018 | -92.9486
9.5021 | -92.9493
9.5025 | -92.9497
9.5028 | -92.9495
9.5031 | -92.9496
9.5034 | -92.9502
9.5037 | -92.9507
9.504 | -92.9513
9.5043 | -92.952
9.5046 | -92.9526
9.5049 | -92.9534
9.5053 | -92.9538
9.5056 | -92.9536
9.5059 | -92.953
9.5062 | -92.9519
9.5065 | -92.9523
9.5068 | -92.9528
9.5071 | -92.9535
9.5074 | -92.954
9.5077 | -92.9552
9.508 | -92.9557
9.5084 | -92.9567
9.5087 | -92.9577
9.509 | -92.9582
9.5093 | -92.9584
9.5096 | -92.9587
9.5099 | -92.9592
9.5102 | -92.9597
9.5105 | -92.9603
9.5108 | -92.9606
9.5112 | -92.961
9.5115 | -92.9612
9.5118 | -92.9616
9.5121 | -92.9613
9.5124 | -92.9617
9.5127 | -92.9622
9.513 | -92.9628
9.5133 | -92.9629
9.5136 | -92.9634
9.5139 | -92.9639
9.5143 | -92.9642
9.5146 | -92.9644
9.5149 | -92.9646
9.5152 | -92.9646
9.5155 | -92.9652
9.5158 | -92.9661
9.5161 | -92.9665
9.5164 | -92.9666
9.5167 | -92.967
9.5171 | -92.9676
9.5174 | -92.9683
9.5177 | -92.9692
9.518 | -92.97
9.5183 | -92.9707
9.5186 | -92.9714
9.5189 | -92.9723
9.5192 | -92.9731
9.5195 | -92.9738
9.5198 | -92.9742
9.5202 | -92.9745
9.5205 | -92.975
9.5208 | -92.9757
9.5211 | -92.9763
9.5214 | -92.9766
9.5217 | -92.9765
9.522 | -92.9768
9.5223 | -92.9771
9.5226 | -92.9774
9.523 | -92.9771
9.5233 | -92.9779
9.5236 | -92.9791
9.5239 | -92.9798
9.5242 | -92.9798
9.5245 | -92.9801
9.5248 | -92.9806
9.5251 | -92.9819
9.5254 | -92.9829
9.5257 | -92.9833
9.5261 | -92.9837
9.5264 | -92.9842
9.5267 | -92.9849
9.527 | -92.9852
9.5273 | -92.9855
9.5276 | -92.9862
9.5279 | -92.9873
9.5282 | -92.988
9.5285 | -92.9885
9.5289 | -92.9898
9.5292 | -92.9906
9.5295 | -92.9912
9.5298 | -92.9922
9.5301 | -92.9933
9.5304 | -92.994
9.5307 | -92.995
9.531 | -92.9961
9.5313 | -92.9965
9.5316 | -92.9969
9.532 | -92.9974
9.5323 | -92.9975
9.5326 | -92.9978
9.5329 | -92.9987
9.5332 | -92.999
9.5335 | -92.9993
9.5338 | -92.9996
9.5341 | -93
9.5344 | -93.0008
9.5348 | -93.0018
9.5351 | -93.0027
9.5354 | -93.0031
9.5357 | -93.0036
9.536 | -93.0037
9.5363 | -93.0038
9.5366 | -93.0039
9.5369 | -93.0046
9.5372 | -93.0052
9.5375 | -93.0054
9.5379 | -93.0054
9.5382 | -93.0063
9.5385 | -93.0074
9.5388 | -93.008
9.5391 | -93.0083
9.5394 | -93.009
9.5397 | -93.0101
9.54 | -93.0113
9.5403 | -93.012
9.5407 | -93.0129
9.541 | -93.0132
9.5413 | -93.0134
9.5416 | -93.0138
9.5419 | -93.0144
9.5422 | -93.015
9.5425 | -93.0154
9.5428 | -93.0157
9.5431 | -93.0157
9.5434 | -93.0161
9.5438 | -93.0165
9.5441 | -93.0168
9.5444 | -93.0174
9.5447 | -93.0185
9.545 | -93.0192
9.5453 | -93.0191
9.5456 | -93.0188
9.5459 | -93.0189
9.5462 | -93.0197
9.5466 | -93.0207
9.5469 | -93.0211
9.5472 | -93.0213
9.5475 | -93.021
9.5478 | -93.0205
9.5481 | -93.0206
9.5484 | -93.0215
9.5487 | -93.022
9.549 | -93.022
9.5493 | -93.0224
9.5497 | -93.0228
9.55 | -93.0237
9.5503 | -93.0242
9.5506 | -93.0243
9.5509 | -93.0248
9.5512 | -93.0252
9.5515 | -93.0259
9.5518 | -93.0268
9.5521 | -93.0273
9.5525 | -93.0277
9.5528 | -93.0283
9.5531 | -93.029
9.5534 | -93.0297
9.5537 | -93.0305
9.554 | -93.0314
9.5543 | -93.0317
9.5546 | -93.0321
9.5549 | -93.0326
9.5552 | -93.0334
9.5556 | -93.0339
9.5559 | -93.0343
9.5562 | -93.0351
9.5565 | -93.0355
9.5568 | -93.0356
9.5571 | -93.0356
9.5574 | -93.0356
9.5577 | -93.0359
9.558 | -93.0367
9.5584 | -93.0371
9.5587 | -93.0373
9.559 | -93.0378
9.5593 | -93.038
9.5596 | -93.0379
9.5599 | -93.0381
9.5602 | -93.0385
9.5605 | -93.0388
9.5608 | -93.0388
9.5611 | -93.0392
9.5615 | -93.0398
9.5618 | -93.0401
9.5621 | -93.0407
9.5624 | -93.0419
9.5627 | -93.0426
9.563 | -93.0431
9.5633 | -93.0441
9.5636 | -93.0453
9.5639 | -93.0462
9.5643 | -93.0465
9.5646 | -93.0468
9.5649 | -93.0467
9.5652 | -93.0471
9.5655 | -93.0476
9.5658 | -93.0481
9.5661 | -93.049
9.5664 | -93.0498
9.5667 | -93.0497
9.567 | -93.0497
9.5674 | -93.0502
9.5677 | -93.0513
9.568 | -93.0521
9.5683 | -93.0531
9.5686 | -93.0536
9.5689 | -93.0543
9.5692 | -93.0552
9.5695 | -93.056
9.5698 | -93.056
9.5702 | -93.0555
9.5705 | -93.0555
9.5708 | -93.0554
9.5711 | -93.0558
9.5714 | -93.0571
9.5717 | -93.0578
9.572 | -93.0581
9.5723 | -93.0584
9.5726 | -93.0579
9.5729 | -93.0574
9.5733 | -93.0581
9.5736 | -93.0589
9.5739 | -93.0597
9.5742 | -93.0598
9.5745 | -93.06
9.5748 | -93.0606
9.5751 | -93.0613
9.5754 | -93.0618
9.5757 | -93.0622
9.5761 | -93.0627
9.5764 | -93.063
9.5767 | -93.0635
9.577 | -93.0641
9.5773 | -93.0644
9.5776 | -93.0649
9.5779 | -93.0653
9.5782 | -93.0659
9.5785 | -93.0659
9.5788 | -93.0658
9.5792 | -93.066
9.5795 | -93.0667
9.5798 | -93.067
9.5801 | -93.067
9.5804 | -93.0678
9.5807 | -93.0685
9.581 | -93.0687
9.5813 | -93.069
9.5816 | -93.0696
9.582 | -93.0697
9.5823 | -93.07
9.5826 | -93.0701
9.5829 | -93.07
9.5832 | -93.0701
9.5835 | -93.0703
9.5838 | -93.0703
9.5841 | -93.0708
9.5844 | -93.0712
9.5847 | -93.0713
9.5851 | -93.0717
9.5854 | -93.072
9.5857 | -93.0725
9.586 | -93.073
9.5863 | -93.074
9.5866 | -93.0748
9.5869 | -93.0755
9.5872 | -93.0758
9.5875 | -93.0761
9.5879 | -93.076
9.5882 | -93.076
9.5885 | -93.0766
9.5888 | -93.077
9.5891 | -93.0777
9.5894 | -93.078
9.5897 | -93.0785
9.59 | -93.0787
9.5903 | -93.079
9.5906 | -93.0797
9.591 | -93.0802
9.5913 | -93.0807
9.5916 | -93.0807
9.5919 | -93.0808
9.5922 | -93.0814
9.5925 | -93.082
9.5928 | -93.0827
9.5931 | -93.0835
9.5934 | -93.0832
9.5938 | -93.0837
9.5941 | -93.0852
9.5944 | -93.0856
9.5947 | -93.0861
9.595 | -93.0867
9.5953 | -93.0874
9.5956 | -93.0866
9.5959 | -93.0873
9.5962 | -93.088
9.5965 | -93.0885
9.5969 | -93.0889
9.5972 | -93.0894
9.5975 | -93.09
9.5978 | -93.0904
9.5981 | -93.091
9.5984 | -93.0913
9.5987 | -93.0916
9.599 | -93.0918
9.5993 | -93.0922
9.5997 | -93.0925
9.6 | -93.0927
9.6003 | -93.0933
9.6006 | -93.0939
9.6009 | -93.0939
9.6012 | -93.0943
9.6015 | -93.0948
9.6018 | -93.0952
9.6021 | -93.0957
9.6024 | -93.0963
9.6028 | -93.0968
9.6031 | -93.0968
9.6034 | -93.0971
9.6037 | -93.0975
9.604 | -93.0979
9.6043 | -93.0983
9.6046 | -93.0989
9.6049 | -93.0993
9.6052 | -93.0995
9.6056 | -93.0997
9.6059 | -93.0995
9.6062 | -93.0994
9.6065 | -93.0995
9.6068 | -93.0999
9.6071 | -93.1003
9.6074 | -93.1008
9.6077 | -93.1015
9.608 | -93.1021
9.6083 | -93.1026
9.6087 | -93.1027
9.609 | -93.1036
9.6093 | -93.1041
9.6096 | -93.1044
9.6099 | -93.104
9.6102 | -93.1037
9.6105 | -93.1035
9.6108 | -93.104
9.6111 | -93.1043
9.6115 | -93.1046
9.6118 | -93.1046
9.6121 | -93.1037
9.6124 | -93.1035
9.6127 | -93.1038
9.613 | -93.1043
9.6133 | -93.105
9.6136 | -93.105
9.6139 | -93.1048
9.6143 | -93.1057
9.6146 | -93.1066
9.6149 | -93.1076
9.6152 | -93.1081
9.6155 | -93.108
9.6158 | -93.1083
9.6161 | -93.1091
9.6164 | -93.1101
9.6167 | -93.1116
9.617 | -93.1123
9.6174 | -93.1124
9.6177 | -93.1121
9.618 | -93.1121
9.6183 | -93.1124
9.6186 | -93.1126
9.6189 | -93.1122
9.6192 | -93.1118
9.6195 | -93.1118
9.6198 | -93.1121
9.6202 | -93.1126
9.6205 | -93.1132
9.6208 | -93.1134
9.6211 | -93.1135
9.6214 | -93.1137
9.6217 | -93.1141
9.622 | -93.1144
9.6223 | -93.1147
9.6226 | -93.1151
9.6229 | -93.1154
9.6233 | -93.1162
9.6236 | -93.1167
9.6239 | -93.1173
9.6242 | -93.1177
9.6245 | -93.1183
9.6248 | -93.1185
9.6251 | -93.1191
9.6254 | -93.1206
9.6257 | -93.1215
9.626 | -93.1222
9.6264 | -93.1226
9.6267 | -93.1229
9.627 | -93.1234
9.6273 | -93.1235
9.6276 | -93.1238
9.6279 | -93.1245
9.6282 | -93.1252
9.6285 | -93.1256
9.6288 | -93.1263
9.6292 | -93.1268
9.6295 | -93.127
9.6298 | -93.1272
9.6301 | -93.1275
9.6304 | -93.1277
9.6307 | -93.1283
9.631 | -93.1293
9.6313 | -93.1301
9.6316 | -93.1297
9.632 | -93.1299
9.6323 | -93.13
9.6326 | -93.1303
9.6329 | -93.1303
9.6332 | -93.1307
9.6335 | -93.1312
9.6338 | -93.1318
9.6341 | -93.1329
9.6344 | -93.1327
9.6347 | -93.1328
9.6351 | -93.1331
9.6354 | -93.133
9.6357 | -93.1332
9.636 | -93.1333
9.6363 | -93.1337
9.6366 | -93.1339
9.6369 | -93.134
9.6372 | -93.1342
9.6375 | -93.1348
9.6379 | -93.1352
9.6382 | -93.1355
9.6385 | -93.1363
9.6388 | -93.1369
9.6391 | -93.1374
9.6394 | -93.1375
9.6397 | -93.1374
9.64 | -93.1376
9.6403 | -93.1376
9.6406 | -93.1379
9.641 | -93.1379
9.6413 | -93.1379
9.6416 | -93.1379
9.6419 | -93.1384
9.6422 | -93.1385
9.6425 | -93.1388
9.6428 | -93.1392
9.6431 | -93.1394
9.6434 | -93.1397
9.6438 | -93.1397
9.6441 | -93.1401
9.6444 | -93.1403
9.6447 | -93.1408
9.645 | -93.1414
9.6453 | -93.1415
9.6456 | -93.141
9.6459 | -93.1402
9.6462 | -93.1397
9.6465 | -93.1394
9.6469 | -93.1396
9.6472 | -93.1399
9.6475 | -93.1401
9.6478 | -93.1405
9.6481 | -93.1409
9.6484 | -93.1411
9.6487 | -93.1418
9.649 | -93.1433
9.6493 | -93.1445
9.6497 | -93.145
9.65 | -93.1451
9.6503 | -93.1454
9.6506 | -93.1457
9.6509 | -93.1461
9.6512 | -93.1465
9.6515 | -93.1464
9.6518 | -93.1485
9.6521 | -93.1497
9.6524 | -93.1499
9.6528 | -93.1506
9.6531 | -93.1511
9.6534 | -93.1513
9.6537 | -93.1514
9.654 | -93.1516
9.6543 | -93.1503
9.6546 | -93.1499
9.6549 | -93.1501
9.6552 | -93.1503
9.6556 | -93.1507
9.6559 | -93.151
9.6562 | -93.1513
9.6565 | -93.1514
9.6568 | -93.1518
9.6571 | -93.152
9.6574 | -93.152
9.6577 | -93.1525
9.658 | -93.1519
9.6583 | -93.1522
9.6587 | -93.1525
9.659 | -93.1525
9.6593 | -93.1523
9.6596 | -93.1524
9.6599 | -93.1527
9.6602 | -93.1531
9.6605 | -93.1534
9.6608 | -93.1536
9.6611 | -93.1537
9.6615 | -93.154
9.6618 | -93.1542
9.6621 | -93.1543
9.6624 | -93.1547
9.6627 | -93.1549
9.663 | -93.1549
9.6633 | -93.155
9.6636 | -93.1552
9.6639 | -93.1554
9.6642 | -93.1557
9.6646 | -93.1559
9.6649 | -93.1559
9.6652 | -93.1562
9.6655 | -93.1564
9.6658 | -93.1568
9.6661 | -93.1576
9.6664 | -93.1577
9.6667 | -93.1579
9.667 | -93.1581
9.6674 | -93.1584
9.6677 | -93.1584
9.668 | -93.1586
9.6683 | -93.1584
9.6686 | -93.1582
9.6689 | -93.1586
9.6692 | -93.1587
9.6695 | -93.159
9.6698 | -93.1593
9.6701 | -93.1602
9.6705 | -93.1607
9.6708 | -93.1609
9.6711 | -93.1612
9.6714 | -93.1613
9.6717 | -93.1613
9.672 | -93.1611
9.6723 | -93.1615
9.6726 | -93.1628
9.6729 | -93.1632
9.6733 | -93.1633
9.6736 | -93.1631
9.6739 | -93.1635
9.6742 | -93.1637
9.6745 | -93.1641
9.6748 | -93.1644
9.6751 | -93.1647
9.6754 | -93.1659
9.6757 | -93.1672
9.676 | -93.1677
9.6764 | -93.168
9.6767 | -93.1678
9.677 | -93.1679
9.6773 | -93.1683
9.6776 | -93.1689
9.6779 | -93.1695
9.6782 | -93.17
9.6785 | -93.1705
9.6788 | -93.171
9.6792 | -93.1713
9.6795 | -93.1716
9.6798 | -93.1718
9.6801 | -93.172
9.6804 | -93.1727
9.6807 | -93.1729
9.681 | -93.1729
9.6813 | -93.1729
9.6816 | -93.1731
9.6819 | -93.1732
9.6823 | -93.1733
9.6826 | -93.1734
9.6829 | -93.1733
9.6832 | -93.1732
9.6835 | -93.1733
9.6838 | -93.1735
9.6841 | -93.1739
9.6844 | -93.1745
9.6847 | -93.1748
9.6851 | -93.1749
9.6854 | -93.175
9.6857 | -93.1756
9.686 | -93.1762
9.6863 | -93.1761
9.6866 | -93.1762
9.6869 | -93.1763
9.6872 | -93.1764
9.6875 | -93.1766
9.6878 | -93.1767
9.6882 | -93.1766
9.6885 | -93.177
9.6888 | -93.1778
9.6891 | -93.1778
9.6894 | -93.1783
9.6897 | -93.1785
9.69 | -93.1786
9.6903 | -93.1796
9.6906 | -93.18
9.691 | -93.1803
9.6913 | -93.1809
9.6916 | -93.1816
9.6919 | -93.1818
9.6922 | -93.1822
9.6925 | -93.1824
9.6928 | -93.1824
9.6931 | -93.1833
9.6934 | -93.1836
9.6937 | -93.1836
9.6941 | -93.1838
9.6944 | -93.1843
9.6947 | -93.1839
9.695 | -93.184
9.6953 | -93.1846
9.6956 | -93.1848
9.6959 | -93.185
9.6962 | -93.1851
9.6965 | -93.1855
9.6969 | -93.1859
9.6972 | -93.1865
9.6975 | -93.1868
9.6978 | -93.187
9.6981 | -93.1869
9.6984 | -93.1869
9.6987 | -93.1869
9.699 | -93.1869
9.6993 | -93.1869
9.6996 | -93.187
9.7 | -93.1869
9.7003 | -93.1867
9.7006 | -93.1857
9.7009 | -93.1851
9.7012 | -93.1854
9.7015 | -93.1857
9.7018 | -93.186
9.7021 | -93.1862
9.7024 | -93.1865
9.7028 | -93.1866
9.7031 | -93.1867
9.7034 | -93.1858
9.7037 | -93.1859
9.704 | -93.1861
9.7043 | -93.1862
9.7046 | -93.1857
9.7049 | -93.1852
9.7052 | -93.1852
9.7055 | -93.1851
9.7059 | -93.1852
9.7062 | -93.1859
9.7065 | -93.1864
9.7068 | -93.1864
9.7071 | -93.187
9.7074 | -93.1869
9.7077 | -93.1869
9.708 | -93.1869
9.7083 | -93.187
9.7087 | -93.1871
9.709 | -93.1873
9.7093 | -93.1874
9.7096 | -93.1877
9.7099 | -93.1877
9.7102 | -93.1875
9.7105 | -93.1872
9.7108 | -93.1872
9.7111 | -93.1874
9.7114 | -93.1875
9.7118 | -93.1875
9.7121 | -93.1876
9.7124 | -93.187
9.7127 | -93.1873
9.713 | -93.1876
9.7133 | -93.1878
9.7136 | -93.1885
9.7139 | -93.1899
9.7142 | -93.1905
9.7146 | -93.1912
9.7149 | -93.1918
9.7152 | -93.1919
9.7155 | -93.1919
9.7158 | -93.1914
9.7161 | -93.1913
9.7164 | -93.192
9.7167 | -93.1923
9.717 | -93.1925
9.7173 | -93.1926
9.7177 | -93.1931
9.718 | -93.1934
9.7183 | -93.1936
9.7186 | -93.1937
9.7189 | -93.194
9.7192 | -93.1946
9.7195 | -93.1949
9.7198 | -93.1956
9.7201 | -93.1969
9.7205 | -93.1976
9.7208 | -93.1979
9.7211 | -93.1981
9.7214 | -93.1983
9.7217 | -93.1983
9.722 | -93.1986
9.7223 | -93.1988
9.7226 | -93.1985
9.7229 | -93.1986
9.7232 | -93.1987
9.7236 | -93.1981
9.7239 | -93.1974
9.7242 | -93.1971
9.7245 | -93.1974
9.7248 | -93.1978
9.7251 | -93.1975
9.7254 | -93.1974
9.7257 | -93.1976
9.726 | -93.1977
9.7264 | -93.1977
9.7267 | -93.1976
9.727 | -93.1975
9.7273 | -93.1975
9.7276 | -93.1977
9.7279 | -93.1979
9.7282 | -93.1981
9.7285 | -93.1982
9.7288 | -93.1983
9.7291 | -93.1982
9.7295 | -93.1981
9.7298 | -93.1984
9.7301 | -93.1982
9.7304 | -93.1987
9.7307 | -93.1986
9.731 | -93.1986
9.7313 | -93.1987
9.7316 | -93.199
9.7319 | -93.1993
9.7323 | -93.1998
9.7326 | -93.2005
9.7329 | -93.2006
9.7332 | -93.2006
9.7335 | -93.2006
9.7338 | -93.2004
9.7341 | -93.2
9.7344 | -93.2007
9.7347 | -93.2015
9.735 | -93.2018
9.7354 | -93.2019
9.7357 | -93.2019
9.736 | -93.2019
9.7363 | -93.2018
9.7366 | -93.2021
9.7369 | -93.2023
9.7372 | -93.2027
9.7375 | -93.2028
9.7378 | -93.2029
9.7382 | -93.2029
9.7385 | -93.2029
9.7388 | -93.2031
9.7391 | -93.2032
9.7394 | -93.2033
9.7397 | -93.2038
9.74 | -93.2042
9.7403 | -93.2047
9.7406 | -93.2048
9.7409 | -93.2049
9.7413 | -93.2052
9.7416 | -93.2051
9.7419 | -93.2047
9.7422 | -93.2046
9.7425 | -93.2048
9.7428 | -93.2054
9.7431 | -93.2057
9.7434 | -93.2058
9.7437 | -93.206
9.7441 | -93.2064
9.7444 | -93.2068
9.7447 | -93.2064
9.745 | -93.2062
9.7453 | -93.2062
9.7456 | -93.2063
9.7459 | -93.2069
9.7462 | -93.2076
9.7465 | -93.2075
9.7469 | -93.2075
9.7472 | -93.2077
9.7475 | -93.2078
9.7478 | -93.208
9.7481 | -93.2083
9.7484 | -93.2083
9.7487 | -93.2086
9.749 | -93.2092
9.7493 | -93.2094
9.7496 | -93.2093
9.75 | -93.2093
9.7503 | -93.2091
9.7506 | -93.2089
9.7509 | -93.2083
9.7512 | -93.208
9.7515 | -93.2078
9.7518 | -93.2069
9.7521 | -93.2065
9.7524 | -93.2064
9.7528 | -93.2058
9.7531 | -93.2059
9.7534 | -93.206
9.7537 | -93.2065
9.754 | -93.207
9.7543 | -93.207
9.7546 | -93.2069
9.7549 | -93.2065
9.7552 | -93.2063
9.7555 | -93.2065
9.7559 | -93.2062
9.7562 | -93.2061
9.7565 | -93.2058
9.7568 | -93.2056
9.7571 | -93.206
9.7574 | -93.2059
9.7577 | -93.2059
9.758 | -93.2062
9.7583 | -93.2067
9.7586 | -93.2067
9.759 | -93.2069
9.7593 | -93.207
9.7596 | -93.207
9.7599 | -93.2071
9.7602 | -93.2073
9.7605 | -93.2077
9.7608 | -93.2077
9.7611 | -93.2076
9.7614 | -93.2078
9.7618 | -93.208
9.7621 | -93.2082
9.7624 | -93.2084
9.7627 | -93.208
9.763 | -93.2078
9.7633 | -93.208
9.7636 | -93.2084
9.7639 | -93.2086
9.7642 | -93.2087
9.7645 | -93.2091
9.7649 | -93.2096
9.7652 | -93.2095
9.7655 | -93.2086
9.7658 | -93.2084
9.7661 | -93.2088
9.7664 | -93.2095
9.7667 | -93.2104
9.767 | -93.2104
9.7673 | -93.2108
9.7677 | -93.2111
9.768 | -93.2111
9.7683 | -93.2108
9.7686 | -93.2103
9.7689 | -93.2108
9.7692 | -93.2116
9.7695 | -93.2118
9.7698 | -93.212
9.7701 | -93.212
9.7705 | -93.2122
9.7708 | -93.2122
9.7711 | -93.2119
9.7714 | -93.2113
9.7717 | -93.2113
9.772 | -93.2113
9.7723 | -93.2113
9.7726 | -93.2114
9.7729 | -93.2114
9.7732 | -93.2113
9.7736 | -93.2107
9.7739 | -93.2093
9.7742 | -93.2085
9.7745 | -93.2085
9.7748 | -93.2088
9.7751 | -93.2092
9.7754 | -93.2095
9.7757 | -93.21
9.776 | -93.2102
9.7764 | -93.2104
9.7767 | -93.2107
9.777 | -93.2111
9.7773 | -93.2119
9.7776 | -93.2124
9.7779 | -93.2128
9.7782 | -93.2131
9.7785 | -93.2133
9.7788 | -93.2139
9.7791 | -93.2145
9.7795 | -93.2151
9.7798 | -93.2153
9.7801 | -93.2154
9.7804 | -93.2154
9.7807 | -93.2153
9.781 | -93.2155
9.7813 | -93.2156
9.7816 | -93.2159
9.7819 | -93.2167
9.7823 | -93.2171
9.7826 | -93.2174
9.7829 | -93.2176
9.7832 | -93.2178
9.7835 | -93.2179
9.7838 | -93.2179
9.7841 | -93.2179
9.7844 | -93.218
9.7847 | -93.218
9.785 | -93.2174
9.7854 | -93.217
9.7857 | -93.217
9.786 | -93.2172
9.7863 | -93.2175
9.7866 | -93.2173
9.7869 | -93.2175
9.7872 | -93.2177
9.7875 | -93.2178
9.7878 | -93.2175
9.7882 | -93.2168
9.7885 | -93.2169
9.7888 | -93.2169
9.7891 | -93.2169
9.7894 | -93.217
9.7897 | -93.2171
9.79 | -93.2178
9.7903 | -93.2189
9.7906 | -93.2193
9.7909 | -93.2192
9.7913 | -93.2191
9.7916 | -93.2193
9.7919 | -93.2194
9.7922 | -93.2191
9.7925 | -93.219
9.7928 | -93.2193
9.7931 | -93.2195
9.7934 | -93.2195
9.7937 | -93.2194
9.7941 | -93.2197
9.7944 | -93.2191
9.7947 | -93.2187
9.795 | -93.2188
9.7953 | -93.2188
9.7956 | -93.2189
9.7959 | -93.219
9.7962 | -93.2191
9.7965 | -93.2192
9.7968 | -93.2194
9.7972 | -93.22
9.7975 | -93.2208
9.7978 | -93.2211
9.7981 | -93.2211
9.7984 | -93.2211
9.7987 | -93.2211
9.799 | -93.2212
9.7993 | -93.2214
9.7996 | -93.2218
9.8 | -93.2221
9.8003 | -93.2222
9.8006 | -93.2225
9.8009 | -93.2225
9.8012 | -93.2221
9.8015 | -93.2216
9.8018 | -93.2219
9.8021 | -93.2222
9.8024 | -93.2229
9.8027 | -93.2232
9.8031 | -93.2234
9.8034 | -93.2234
9.8037 | -93.2224
9.804 | -93.2214
9.8043 | -93.2212
9.8046 | -93.2213
9.8049 | -93.2213
9.8052 | -93.2213
9.8055 | -93.2217
9.8059 | -93.222
9.8062 | -93.2219
9.8065 | -93.222
9.8068 | -93.2222
9.8071 | -93.2223
9.8074 | -93.2222
9.8077 | -93.2222
9.808 | -93.2222
9.8083 | -93.2222
9.8086 | -93.2226
9.809 | -93.2234
9.8093 | -93.2236
9.8096 | -93.2241
9.8099 | -93.2247
9.8102 | -93.2246
9.8105 | -93.2238
9.8108 | -93.222
9.8111 | -93.2221
9.8114 | -93.2219
9.8118 | -93.2221
9.8121 | -93.2223
9.8124 | -93.2226
9.8127 | -93.2229
9.813 | -93.2235
9.8133 | -93.2239
9.8136 | -93.2244
9.8139 | -93.2247
9.8142 | -93.225
9.8145 | -93.225
9.8149 | -93.2245
9.8152 | -93.2245
9.8155 | -93.2246
9.8158 | -93.2248
9.8161 | -93.2247
9.8164 | -93.2244
9.8167 | -93.2247
9.817 | -93.2242
9.8173 | -93.224
9.8177 | -93.2236
9.818 | -93.2234
9.8183 | -93.2238
9.8186 | -93.2241
9.8189 | -93.2242
9.8192 | -93.2242
9.8195 | -93.2243
9.8198 | -93.2248
9.8201 | -93.2249
9.8204 | -93.2251
9.8208 | -93.2253
9.8211 | -93.2256
9.8214 | -93.2257
9.8217 | -93.2256
9.822 | -93.2258
9.8223 | -93.2261
9.8226 | -93.2261
9.8229 | -93.226
9.8232 | -93.2262
9.8236 | -93.2267
9.8239 | -93.2267
9.8242 | -93.2263
9.8245 | -93.2263
9.8248 | -93.2263
9.8251 | -93.2264
9.8254 | -93.2264
9.8257 | -93.2267
9.826 | -93.227
9.8263 | -93.2269
9.8267 | -93.2269
9.827 | -93.2267
9.8273 | -93.2272
9.8276 | -93.2271
9.8279 | -93.227
9.8282 | -93.2273
9.8285 | -93.2274
9.8288 | -93.228
9.8291 | -93.2283
9.8295 | -93.2283
9.8298 | -93.2285
9.8301 | -93.2289
9.8304 | -93.2293
9.8307 | -93.2291
9.831 | -93.2289
9.8313 | -93.2288
9.8316 | -93.2291
9.8319 | -93.2289
9.8322 | -93.2289
9.8326 | -93.2286
9.8329 | -93.2286
9.8332 | -93.2287
9.8335 | -93.2291
9.8338 | -93.2295
9.8341 | -93.23
9.8344 | -93.2303
9.8347 | -93.2304
9.835 | -93.2308
9.8354 | -93.2313
9.8357 | -93.2315
9.836 | -93.2326
9.8363 | -93.2333
9.8366 | -93.2334
9.8369 | -93.2338
9.8372 | -93.2344
9.8375 | -93.2349
9.8378 | -93.2351
9.8381 | -93.235
9.8385 | -93.2349
9.8388 | -93.235
9.8391 | -93.2351
9.8394 | -93.2349
9.8397 | -93.2348
9.84 | -93.2348
9.8403 | -93.2351
9.8406 | -93.2353
9.8409 | -93.2352
9.8413 | -93.2352
9.8416 | -93.2354
9.8419 | -93.2358
9.8422 | -93.2363
9.8425 | -93.2371
9.8428 | -93.2374
9.8431 | -93.2377
9.8434 | -93.238
9.8437 | -93.2383
9.844 | -93.239
9.8444 | -93.2389
9.8447 | -93.2387
9.845 | -93.2389
9.8453 | -93.2392
9.8456 | -93.2407
9.8459 | -93.2418
9.8462 | -93.242
9.8465 | -93.2423
9.8468 | -93.2427
9.8472 | -93.2426
9.8475 | -93.2429
9.8478 | -93.243
9.8481 | -93.2432
9.8484 | -93.2432
9.8487 | -93.2431
9.849 | -93.2432
9.8493 | -93.2433
9.8496 | -93.2437
9.8499 | -93.2439
9.8503 | -93.2441
9.8506 | -93.2443
9.8509 | -93.2443
9.8512 | -93.2443
9.8515 | -93.2445
9.8518 | -93.2446
9.8521 | -93.2447
9.8524 | -93.2449
9.8527 | -93.2453
9.8531 | -93.2455
9.8534 | -93.2459
9.8537 | -93.2466
9.854 | -93.2468
9.8543 | -93.2471
9.8546 | -93.2476
9.8549 | -93.248
9.8552 | -93.2487
9.8555 | -93.249
9.8558 | -93.2499
9.8562 | -93.2502
9.8565 | -93.2506
9.8568 | -93.2513
9.8571 | -93.2516
9.8574 | -93.2517
9.8577 | -93.2516
9.858 | -93.2515
9.8583 | -93.2516
9.8586 | -93.2517
9.859 | -93.2519
9.8593 | -93.2521
9.8596 | -93.252
9.8599 | -93.2518
9.8602 | -93.2518
9.8605 | -93.2516
9.8608 | -93.2517
9.8611 | -93.2518
9.8614 | -93.2519
9.8617 | -93.2521
9.8621 | -93.2522
9.8624 | -93.2524
9.8627 | -93.2523
9.863 | -93.2522
9.8633 | -93.2523
9.8636 | -93.2523
9.8639 | -93.2527
9.8642 | -93.2529
9.8645 | -93.2534
9.8649 | -93.2532
9.8652 | -93.2528
9.8655 | -93.2529
9.8658 | -93.2527
9.8661 | -93.2524
9.8664 | -93.2523
9.8667 | -93.2524
9.867 | -93.2526
9.8673 | -93.2528
9.8676 | -93.2534
9.868 | -93.2535
9.8683 | -93.2536
9.8686 | -93.2541
9.8689 | -93.255
9.8692 | -93.2551
9.8695 | -93.2551
9.8698 | -93.2553
9.8701 | -93.2555
9.8704 | -93.2556
9.8708 | -93.2555
9.8711 | -93.2556
9.8714 | -93.2559
9.8717 | -93.2559
9.872 | -93.2561
9.8723 | -93.257
9.8726 | -93.2575
9.8729 | -93.2574
9.8732 | -93.2568
9.8735 | -93.2568
9.8739 | -93.257
9.8742 | -93.2573
9.8745 | -93.2574
9.8748 | -93.2574
9.8751 | -93.2575
9.8754 | -93.2577
9.8757 | -93.2573
9.876 | -93.2572
9.8763 | -93.2573
9.8767 | -93.2575
9.877 | -93.2577
9.8773 | -93.2583
9.8776 | -93.2585
9.8779 | -93.2588
9.8782 | -93.2589
9.8785 | -93.2591
9.8788 | -93.2587
9.8791 | -93.2585
9.8794 | -93.2588
9.8798 | -93.2592
9.8801 | -93.2595
9.8804 | -93.2598
9.8807 | -93.2603
9.881 | -93.2605
9.8813 | -93.2608
9.8816 | -93.261
9.8819 | -93.2613
9.8822 | -93.261
9.8826 | -93.2602
9.8829 | -93.2603
9.8832 | -93.2617
9.8835 | -93.2626
9.8838 | -93.262
9.8841 | -93.2615
9.8844 | -93.2612
9.8847 | -93.2613
9.885 | -93.2614
9.8853 | -93.2615
9.8857 | -93.2617
9.886 | -93.2623
9.8863 | -93.2627
9.8866 | -93.2628
9.8869 | -93.2624
9.8872 | -93.2626
9.8875 | -93.2628
9.8878 | -93.263
9.8881 | -93.2632
9.8885 | -93.2636
9.8888 | -93.2639
9.8891 | -93.2642
9.8894 | -93.2641
9.8897 | -93.2641
9.89 | -93.2644
9.8903 | -93.2647
9.8906 | -93.2649
9.8909 | -93.2653
9.8912 | -93.2659
9.8916 | -93.2664
9.8919 | -93.2666
9.8922 | -93.2665
9.8925 | -93.2662
9.8928 | -93.2667
9.8931 | -93.267
9.8934 | -93.2675
9.8937 | -93.2678
9.894 | -93.2678
9.8944 | -93.2682
9.8947 | -93.2684
9.895 | -93.2674
9.8953 | -93.2665
9.8956 | -93.266
9.8959 | -93.2661
9.8962 | -93.2662
9.8965 | -93.2668
9.8968 | -93.267
9.8971 | -93.2669
9.8975 | -93.2669
9.8978 | -93.2676
9.8981 | -93.2681
9.8984 | -93.2687
9.8987 | -93.2692
9.899 | -93.2696
9.8993 | -93.2697
9.8996 | -93.2701
9.8999 | -93.2703
9.9003 | -93.271
9.9006 | -93.2712
9.9009 | -93.2715
9.9012 | -93.2713
9.9015 | -93.2715
9.9018 | -93.2721
9.9021 | -93.2725
9.9024 | -93.2727
9.9027 | -93.2732
9.9031 | -93.2735
9.9034 | -93.2738
9.9037 | -93.2743
9.904 | -93.2747
9.9043 | -93.2747
9.9046 | -93.2743
9.9049 | -93.2743
9.9052 | -93.2748
9.9055 | -93.2753
9.9058 | -93.2755
9.9062 | -93.2755
9.9065 | -93.2757
9.9068 | -93.2761
9.9071 | -93.2754
9.9074 | -93.2748
9.9077 | -93.2747
9.908 | -93.2748
9.9083 | -93.2752
9.9086 | -93.2754
9.909 | -93.2756
9.9093 | -93.2761
9.9096 | -93.2767
9.9099 | -93.2777
9.9102 | -93.2784
9.9105 | -93.2785
9.9108 | -93.2788
9.9111 | -93.2788
9.9114 | -93.2789
9.9117 | -93.2796
9.9121 | -93.2799
9.9124 | -93.2801
9.9127 | -93.2803
9.913 | -93.2805
9.9133 | -93.2806
9.9136 | -93.2808
9.9139 | -93.2809
9.9142 | -93.281
9.9145 | -93.2815
9.9149 | -93.2819
9.9152 | -93.282
9.9155 | -93.2818
9.9158 | -93.2819
9.9161 | -93.2818
9.9164 | -93.2821
9.9167 | -93.2824
9.917 | -93.2832
9.9173 | -93.2845
9.9176 | -93.2856
9.918 | -93.2862
9.9183 | -93.2865
9.9186 | -93.2866
9.9189 | -93.2868
9.9192 | -93.2869
9.9195 | -93.287
9.9198 | -93.2872
9.9201 | -93.2872
9.9204 | -93.2871
9.9208 | -93.287
9.9211 | -93.2871
9.9214 | -93.2875
9.9217 | -93.2878
9.922 | -93.2878
9.9223 | -93.2877
9.9226 | -93.2881
9.9229 | -93.2888
9.9232 | -93.2891
9.9235 | -93.2893
9.9239 | -93.2897
9.9242 | -93.2903
9.9245 | -93.2906
9.9248 | -93.2908
9.9251 | -93.2912
9.9254 | -93.292
9.9257 | -93.2923
9.926 | -93.2923
9.9263 | -93.2914
9.9267 | -93.2916
9.927 | -93.2914
9.9273 | -93.292
9.9276 | -93.2919
9.9279 | -93.2922
9.9282 | -93.292
9.9285 | -93.2917
9.9288 | -93.292
9.9291 | -93.2921
9.9294 | -93.2925
9.9298 | -93.2917
9.9301 | -93.2915
9.9304 | -93.2919
9.9307 | -93.2922
9.931 | -93.2922
9.9313 | -93.2927
9.9316 | -93.2934
9.9319 | -93.2935
9.9322 | -93.2934
9.9326 | -93.2935
9.9329 | -93.2939
9.9332 | -93.2942
9.9335 | -93.2942
9.9338 | -93.2942
9.9341 | -93.2941
9.9344 | -93.2947
9.9347 | -93.2952
9.935 | -93.295
9.9353 | -93.295
9.9357 | -93.2952
9.936 | -93.2951
9.9363 | -93.2954
9.9366 | -93.2955
9.9369 | -93.2962
9.9372 | -93.2958
9.9375 | -93.296
9.9378 | -93.2967
9.9381 | -93.2969
9.9385 | -93.2972
9.9388 | -93.2976
9.9391 | -93.2978
9.9394 | -93.2979
9.9397 | -93.2981
9.94 | -93.2985
9.9403 | -93.2991
9.9406 | -93.2992
9.9409 | -93.2995
9.9412 | -93.2996
9.9416 | -93.2996
9.9419 | -93.2999
9.9422 | -93.3
9.9425 | -93.3001
9.9428 | -93.3004
9.9431 | -93.3005
9.9434 | -93.3005
9.9437 | -93.3002
9.944 | -93.3004
9.9444 | -93.3007
9.9447 | -93.3015
9.945 | -93.3021
9.9453 | -93.3023
9.9456 | -93.3025
9.9459 | -93.3025
9.9462 | -93.3018
9.9465 | -93.3013
9.9468 | -93.3013
9.9471 | -93.3013
9.9475 | -93.3016
9.9478 | -93.3018
9.9481 | -93.3019
9.9484 | -93.3042
9.9487 | -93.3088
9.949 | -93.3099
9.9493 | -93.3103
9.9496 | -93.3103
9.9499 | -93.3104
9.9503 | -93.3106
9.9506 | -93.3108
9.9509 | -93.3111
9.9512 | -93.3111
9.9515 | -93.3115
9.9518 | -93.312
9.9521 | -93.3122
9.9524 | -93.3129
9.9527 | -93.3126
9.953 | -93.3122
9.9534 | -93.312
9.9537 | -93.3121
9.954 | -93.312
9.9543 | -93.3121
9.9546 | -93.3126
9.9549 | -93.314
9.9552 | -93.3152
9.9555 | -93.3155
9.9558 | -93.3159
9.9562 | -93.3158
9.9565 | -93.3159
9.9568 | -93.3163
9.9571 | -93.3165
9.9574 | -93.3166
9.9577 | -93.3163
9.958 | -93.3161
9.9583 | -93.3165
9.9586 | -93.3165
9.9589 | -93.3164
9.9593 | -93.3166
9.9596 | -93.3167
9.9599 | -93.3169
9.9602 | -93.317
9.9605 | -93.3172
9.9608 | -93.3177
9.9611 | -93.3182
9.9614 | -93.3192
9.9617 | -93.3191
9.9621 | -93.3189
9.9624 | -93.319
9.9627 | -93.3191
9.963 | -93.3193
9.9633 | -93.3193
9.9636 | -93.3195
9.9639 | -93.3195
9.9642 | -93.3196
9.9645 | -93.3199
9.9648 | -93.3202
9.9652 | -93.3205
9.9655 | -93.3202
9.9658 | -93.3201
9.9661 | -93.3202
9.9664 | -93.3205
9.9667 | -93.3207
9.967 | -93.3208
9.9673 | -93.321
9.9676 | -93.3211
9.968 | -93.3212
9.9683 | -93.3217
9.9686 | -93.3218
9.9689 | -93.321
9.9692 | -93.3206
9.9695 | -93.3214
9.9698 | -93.3215
9.9701 | -93.3215
9.9704 | -93.3216
9.9707 | -93.3219
9.9711 | -93.3218
9.9714 | -93.322
9.9717 | -93.3225
9.972 | -93.3231
9.9723 | -93.3242
9.9726 | -93.3246
9.9729 | -93.325
9.9732 | -93.3253
9.9735 | -93.3255
9.9739 | -93.3257
9.9742 | -93.3259
9.9745 | -93.3257
9.9748 | -93.3259
9.9751 | -93.3267
9.9754 | -93.3266
9.9757 | -93.3269
9.976 | -93.327
9.9763 | -93.3272
9.9766 | -93.3273
9.977 | -93.3272
9.9773 | -93.3266
9.9776 | -93.327
9.9779 | -93.3272
9.9782 | -93.3273
9.9785 | -93.3276
9.9788 | -93.3278
9.9791 | -93.3278
9.9794 | -93.3275
9.9798 | -93.3271
9.9801 | -93.3271
9.9804 | -93.3274
9.9807 | -93.3274
9.981 | -93.3277
9.9813 | -93.3287
9.9816 | -93.3293
9.9819 | -93.3298
9.9822 | -93.3306
9.9825 | -93.3309
9.9829 | -93.331
9.9832 | -93.3312
9.9835 | -93.3313
9.9838 | -93.3311
9.9841 | -93.3307
9.9844 | -93.3307
9.9847 | -93.3305
9.985 | -93.3307
9.9853 | -93.3306
9.9857 | -93.3307
9.986 | -93.3308
9.9863 | -93.3311
9.9866 | -93.3313
9.9869 | -93.3316
9.9872 | -93.3319
9.9875 | -93.3319
9.9878 | -93.3312
9.9881 | -93.3308
9.9884 | -93.3314
9.9888 | -93.3314
9.9891 | -93.3318
9.9894 | -93.3323
9.9897 | -93.3324
9.99 | -93.3325
9.9903 | -93.3329
9.9906 | -93.3331
9.9909 | -93.333
9.9912 | -93.3331
9.9916 | -93.3334
9.9919 | -93.3331
9.9922 | -93.3332
9.9925 | -93.3335
9.9928 | -93.3336
9.9931 | -93.3332
9.9934 | -93.333
9.9937 | -93.3332
9.994 | -93.3333
9.9943 | -93.3334
9.9947 | -93.3334
9.995 | -93.3337
9.9953 | -93.3339
9.9956 | -93.3343
9.9959 | -93.3344
9.9962 | -93.3344
9.9965 | -93.3357
9.9968 | -93.3369
9.9971 | -93.337
9.9975 | -93.3371
9.9978 | -93.3374
9.9981 | -93.3375
9.9984 | -93.3378
9.9987 | -93.3381
9.999 | -93.3386
9.9993 | -93.3388
9.9996 | -93.339
9.9999 | -93.3392
10.0002 | -93.3396
10.0006 | -93.34
10.0009 | -93.3402
10.0012 | -93.3404
10.0015 | -93.3404
10.0018 | -93.3402
10.0021 | -93.3396
10.0024 | -93.3398
10.0027 | -93.3398
10.003 | -93.3398
10.0034 | -93.3398
10.0037 | -93.3403
10.004 | -93.3407
10.0043 | -93.3407
10.0046 | -93.3409
10.0049 | -93.3412
10.0052 | -93.3415
10.0055 | -93.3415
10.0058 | -93.3414
10.0061 | -93.3419
10.0065 | -93.3425
10.0068 | -93.3428
10.0071 | -93.343
10.0074 | -93.3431
10.0077 | -93.3431
10.008 | -93.3432
10.0083 | -93.3431
10.0086 | -93.3427
10.0089 | -93.3425
10.0093 | -93.3433
10.0096 | -93.3444
10.0099 | -93.3445
10.0102 | -93.3448
10.0105 | -93.3451
10.0108 | -93.3455
10.0111 | -93.3459
10.0114 | -93.3461
10.0117 | -93.3463
10.012 | -93.3462
10.0124 | -93.3467
10.0127 | -93.3467
10.013 | -93.347
10.0133 | -93.3467
10.0136 | -93.3469
10.0139 | -93.347
10.0142 | -93.3473
10.0145 | -93.3473
10.0148 | -93.3473
10.0152 | -93.3472
10.0155 | -93.3472
10.0158 | -93.3476
10.0161 | -93.3477
10.0164 | -93.3476
10.0167 | -93.3479
10.017 | -93.3482
10.0173 | -93.3487
10.0176 | -93.3486
10.0179 | -93.3482
10.0183 | -93.3477
10.0186 | -93.348
10.0189 | -93.3486
10.0192 | -93.3488
10.0195 | -93.3488
10.0198 | -93.349
10.0201 | -93.3491
10.0204 | -93.3493
10.0207 | -93.3494
10.0211 | -93.3492
10.0214 | -93.349
10.0217 | -93.3489
10.022 | -93.3487
10.0223 | -93.3486
10.0226 | -93.3486
10.0229 | -93.3487
10.0232 | -93.3487
10.0235 | -93.3487
10.0238 | -93.3493
10.0242 | -93.35
10.0245 | -93.3498
10.0248 | -93.3498
10.0251 | -93.3502
10.0254 | -93.3501
10.0257 | -93.3493
10.026 | -93.3493
10.0263 | -93.3492
10.0266 | -93.3494
10.027 | -93.3497
10.0273 | -93.35
10.0276 | -93.3502
10.0279 | -93.3504
10.0282 | -93.3505
10.0285 | -93.3506
10.0288 | -93.351
10.0291 | -93.3514
10.0294 | -93.3513
10.0297 | -93.3516
10.0301 | -93.3526
10.0304 | -93.3536
10.0307 | -93.3534
10.031 | -93.3533
10.0313 | -93.3534
10.0316 | -93.3531
10.0319 | -93.3526
10.0322 | -93.3526
10.0325 | -93.3531
10.0329 | -93.3533
10.0332 | -93.3535
10.0335 | -93.3536
10.0338 | -93.3536
10.0341 | -93.3544
10.0344 | -93.3546
10.0347 | -93.354
10.035 | -93.3532
10.0353 | -93.353
10.0357 | -93.3533
10.036 | -93.3541
10.0363 | -93.3548
10.0366 | -93.3557
10.0369 | -93.3568
10.0372 | -93.3571
10.0375 | -93.3571
10.0378 | -93.3572
10.0381 | -93.3572
10.0384 | -93.3568
10.0388 | -93.357
10.0391 | -93.3572
10.0394 | -93.3575
10.0397 | -93.3577
10.04 | -93.3575
10.0403 | -93.3576
10.0406 | -93.3577
10.0409 | -93.3569
10.0412 | -93.3554
10.0416 | -93.3557
10.0419 | -93.3559
10.0422 | -93.3559
10.0425 | -93.3559
10.0428 | -93.3561
10.0431 | -93.356
10.0434 | -93.3561
10.0437 | -93.3564
10.044 | -93.3565
10.0443 | -93.3566
10.0447 | -93.3571
10.045 | -93.3573
10.0453 | -93.3571
10.0456 | -93.357
10.0459 | -93.3573
10.0462 | -93.3577
10.0465 | -93.3575
10.0468 | -93.3568
10.0471 | -93.3566
10.0474 | -93.3569
10.0478 | -93.3573
10.0481 | -93.3575
10.0484 | -93.3575
10.0487 | -93.3576
10.049 | -93.3575
10.0493 | -93.3576
10.0496 | -93.3577
10.0499 | -93.3578
10.0502 | -93.3581
10.0506 | -93.3585
10.0509 | -93.3583
10.0512 | -93.3579
10.0515 | -93.3575
10.0518 | -93.3574
10.0521 | -93.3574
10.0524 | -93.357
10.0527 | -93.3567
10.053 | -93.3566
10.0534 | -93.3568
10.0537 | -93.3569
10.054 | -93.357
10.0543 | -93.3573
10.0546 | -93.3576
10.0549 | -93.3577
10.0552 | -93.3573
10.0555 | -93.3571
10.0558 | -93.357
10.0561 | -93.357
10.0565 | -93.3569
10.0568 | -93.3563
10.0571 | -93.3557
10.0574 | -93.3559
10.0577 | -93.3561
10.058 | -93.3564
10.0583 | -93.3567
10.0586 | -93.3573
10.0589 | -93.3587
10.0593 | -93.3593
10.0596 | -93.3596
10.0599 | -93.3598
10.0602 | -93.3599
10.0605 | -93.3599
10.0608 | -93.3601
10.0611 | -93.3603
10.0614 | -93.3602
10.0617 | -93.3596
10.062 | -93.3593
10.0624 | -93.3596
10.0627 | -93.3599
10.063 | -93.36
10.0633 | -93.36
10.0636 | -93.3599
10.0639 | -93.3598
10.0642 | -93.3604
10.0645 | -93.3611
10.0648 | -93.3611
10.0652 | -93.3605
10.0655 | -93.3599
10.0658 | -93.36
10.0661 | -93.3599
10.0664 | -93.3599
10.0667 | -93.3599
10.067 | -93.3597
10.0673 | -93.3593
10.0676 | -93.3592
10.0679 | -93.3588
10.0683 | -93.3588
10.0686 | -93.3588
10.0689 | -93.3587
10.0692 | -93.3588
10.0695 | -93.3582
10.0698 | -93.3581
10.0701 | -93.3583
10.0704 | -93.3583
10.0707 | -93.3583
10.0711 | -93.3586
10.0714 | -93.3586
10.0717 | -93.3587
10.072 | -93.3586
10.0723 | -93.3585
10.0726 | -93.3585
10.0729 | -93.3586
10.0732 | -93.3587
10.0735 | -93.3589
10.0738 | -93.359
10.0742 | -93.3591
10.0745 | -93.3593
10.0748 | -93.3596
10.0751 | -93.3597
10.0754 | -93.3596
10.0757 | -93.3596
10.076 | -93.3595
10.0763 | -93.359
10.0766 | -93.3591
10.077 | -93.3593
10.0773 | -93.3595
10.0776 | -93.3594
10.0779 | -93.3591
10.0782 | -93.3586
10.0785 | -93.3583
10.0788 | -93.3584
10.0791 | -93.3586
10.0794 | -93.3589
10.0797 | -93.3587
10.0801 | -93.359
10.0804 | -93.359
10.0807 | -93.359
10.081 | -93.3594
10.0813 | -93.3592
10.0816 | -93.3591
10.0819 | -93.3593
10.0822 | -93.3595
10.0825 | -93.3598
10.0829 | -93.3601
10.0832 | -93.3602
10.0835 | -93.3605
10.0838 | -93.3606
10.0841 | -93.3609
10.0844 | -93.3608
10.0847 | -93.3607
10.085 | -93.3607
10.0853 | -93.361
10.0856 | -93.3613
10.086 | -93.3611
10.0863 | -93.3611
10.0866 | -93.3614
10.0869 | -93.3615
10.0872 | -93.3616
10.0875 | -93.3616
10.0878 | -93.3611
10.0881 | -93.3611
10.0884 | -93.361
10.0888 | -93.3609
10.0891 | -93.361
10.0894 | -93.361
10.0897 | -93.361
10.09 | -93.361
10.0903 | -93.3612
10.0906 | -93.3616
10.0909 | -93.3617
10.0912 | -93.3615
10.0915 | -93.3614
10.0919 | -93.3603
10.0922 | -93.3597
10.0925 | -93.3597
10.0928 | -93.3602
10.0931 | -93.3608
10.0934 | -93.3607
10.0937 | -93.3608
10.094 | -93.3609
10.0943 | -93.3609
10.0947 | -93.3614
10.095 | -93.3613
10.0953 | -93.3611
10.0956 | -93.3611
10.0959 | -93.3612
10.0962 | -93.3614
10.0965 | -93.3615
10.0968 | -93.3608
10.0971 | -93.3598
10.0974 | -93.3594
10.0978 | -93.3589
10.0981 | -93.3587
10.0984 | -93.3588
10.0987 | -93.3592
10.099 | -93.3592
10.0993 | -93.3592
10.0996 | -93.3592
10.0999 | -93.3594
10.1002 | -93.3591
10.1006 | -93.3589
10.1009 | -93.3589
10.1012 | -93.359
10.1015 | -93.3595
10.1018 | -93.3597
10.1021 | -93.3596
10.1024 | -93.3594
10.1027 | -93.3596
10.103 | -93.3598
10.1033 | -93.3597
10.1037 | -93.3596
10.104 | -93.359
10.1043 | -93.3589
10.1046 | -93.3589
10.1049 | -93.3584
10.1052 | -93.3583
10.1055 | -93.3579
10.1058 | -93.3577
10.1061 | -93.3575
10.1065 | -93.3576
10.1068 | -93.3582
10.1071 | -93.3582
10.1074 | -93.3577
10.1077 | -93.3573
10.108 | -93.3572
10.1083 | -93.3573
10.1086 | -93.3571
10.1089 | -93.3567
10.1092 | -93.3564
10.1096 | -93.3565
10.1099 | -93.3564
10.1102 | -93.3563
10.1105 | -93.3563
10.1108 | -93.3568
10.1111 | -93.357
10.1114 | -93.357
10.1117 | -93.357
10.112 | -93.3567
10.1124 | -93.3556
10.1127 | -93.3551
10.113 | -93.3546
10.1133 | -93.3547
10.1136 | -93.3546
10.1139 | -93.3546
10.1142 | -93.3547
10.1145 | -93.3546
10.1148 | -93.3546
10.1151 | -93.3548
10.1155 | -93.3547
10.1158 | -93.3546
10.1161 | -93.3547
10.1164 | -93.3547
10.1167 | -93.3544
10.117 | -93.3544
10.1173 | -93.3544
10.1176 | -93.354
10.1179 | -93.3535
10.1183 | -93.3533
10.1186 | -93.3533
10.1189 | -93.353
10.1192 | -93.3528
10.1195 | -93.3529
10.1198 | -93.3532
10.1201 | -93.3533
10.1204 | -93.3533
10.1207 | -93.3533
10.121 | -93.3537
10.1214 | -93.354
10.1217 | -93.3538
10.122 | -93.3536
10.1223 | -93.3534
10.1226 | -93.3532
10.1229 | -93.3531
10.1232 | -93.353
10.1235 | -93.3528
10.1238 | -93.353
10.1242 | -93.3532
10.1245 | -93.3532
10.1248 | -93.3538
10.1251 | -93.3541
10.1254 | -93.354
10.1257 | -93.354
10.126 | -93.3541
10.1263 | -93.3542
10.1266 | -93.3539
10.1269 | -93.3539
10.1273 | -93.3545
10.1276 | -93.3546
10.1279 | -93.3552
10.1282 | -93.3562
10.1285 | -93.3563
10.1288 | -93.3565
10.1291 | -93.3565
10.1294 | -93.3562
10.1297 | -93.3551
10.1301 | -93.3545
10.1304 | -93.3544
10.1307 | -93.3534
10.131 | -93.353
10.1313 | -93.3538
10.1316 | -93.3546
10.1319 | -93.355
10.1322 | -93.3541
10.1325 | -93.3536
10.1328 | -93.3537
10.1332 | -93.3534
10.1335 | -93.3533
10.1338 | -93.3531
10.1341 | -93.3533
10.1344 | -93.3534
10.1347 | -93.3536
10.135 | -93.3537
10.1353 | -93.3535
10.1356 | -93.3534
10.136 | -93.3533
10.1363 | -93.3535
10.1366 | -93.3533
10.1369 | -93.353
10.1372 | -93.353
10.1375 | -93.3528
10.1378 | -93.3527
10.1381 | -93.3525
10.1384 | -93.3523
10.1387 | -93.3521
10.1391 | -93.3521
10.1394 | -93.3522
10.1397 | -93.3522
10.14 | -93.3521
10.1403 | -93.3522
10.1406 | -93.3523
10.1409 | -93.3519
10.1412 | -93.3513
10.1415 | -93.3513
10.1419 | -93.3513
10.1422 | -93.3513
10.1425 | -93.3513
10.1428 | -93.3513
10.1431 | -93.351
10.1434 | -93.3509
10.1437 | -93.3508
10.144 | -93.3508
10.1443 | -93.3507
10.1446 | -93.3508
10.145 | -93.3507
10.1453 | -93.3506
10.1456 | -93.3502
10.1459 | -93.3503
10.1462 | -93.3505
10.1465 | -93.3507
10.1468 | -93.3509
10.1471 | -93.3505
10.1474 | -93.3498
10.1478 | -93.35
10.1481 | -93.35
10.1484 | -93.3491
10.1487 | -93.3486
10.149 | -93.3491
10.1493 | -93.3497
10.1496 | -93.3497
10.1499 | -93.3503
10.1502 | -93.3509
10.1505 | -93.3507
10.1509 | -93.3505
10.1512 | -93.3504
10.1515 | -93.3506
10.1518 | -93.351
10.1521 | -93.3505
10.1524 | -93.35
10.1527 | -93.3502
10.153 | -93.3502
10.1533 | -93.3503
10.1537 | -93.3502
10.154 | -93.3501
10.1543 | -93.3499
10.1546 | -93.35
10.1549 | -93.3502
10.1552 | -93.3503
10.1555 | -93.3495
10.1558 | -93.3494
10.1561 | -93.3495
10.1564 | -93.3494
10.1568 | -93.3491
10.1571 | -93.3489
10.1574 | -93.3489
10.1577 | -93.3488
10.158 | -93.3485
10.1583 | -93.3484
10.1586 | -93.3487
10.1589 | -93.3486
10.1592 | -93.3484
10.1596 | -93.3483
10.1599 | -93.3487
10.1602 | -93.3486
10.1605 | -93.349
10.1608 | -93.3489
10.1611 | -93.3487
10.1614 | -93.3485
10.1617 | -93.3486
10.162 | -93.3485
10.1623 | -93.3487
10.1627 | -93.3486
10.163 | -93.3482
10.1633 | -93.348
10.1636 | -93.3479
10.1639 | -93.3477
10.1642 | -93.3477
10.1645 | -93.3477
10.1648 | -93.3476
10.1651 | -93.3475
10.1655 | -93.3474
10.1658 | -93.3473
10.1661 | -93.3467
10.1664 | -93.346
10.1667 | -93.3459
10.167 | -93.3459
10.1673 | -93.3459
10.1676 | -93.3457
10.1679 | -93.3454
10.1682 | -93.345
10.1686 | -93.3448
10.1689 | -93.3444
10.1692 | -93.3446
10.1695 | -93.3448
10.1698 | -93.3452
10.1701 | -93.3449
10.1704 | -93.3446
10.1707 | -93.3445
10.171 | -93.3444
10.1714 | -93.3444
10.1717 | -93.3443
10.172 | -93.344
10.1723 | -93.3438
10.1726 | -93.3437
10.1729 | -93.3433
10.1732 | -93.3433
10.1735 | -93.3437
10.1738 | -93.3441
10.1741 | -93.3441
10.1745 | -93.3441
10.1748 | -93.3441
10.1751 | -93.3439
10.1754 | -93.3436
10.1757 | -93.343
10.176 | -93.3426
10.1763 | -93.3426
10.1766 | -93.3424
10.1769 | -93.3422
10.1773 | -93.3424
10.1776 | -93.343
10.1779 | -93.3429
10.1782 | -93.3431
10.1785 | -93.3429
10.1788 | -93.3426
10.1791 | -93.342
10.1794 | -93.3416
10.1797 | -93.3414
10.18 | -93.3413
10.1804 | -93.3411
10.1807 | -93.3411
10.181 | -93.3413
10.1813 | -93.3413
10.1816 | -93.3413
10.1819 | -93.3418
10.1822 | -93.3424
10.1825 | -93.3428
10.1828 | -93.3432
10.1832 | -93.3434
10.1835 | -93.3442
10.1838 | -93.3451
10.1841 | -93.3454
10.1844 | -93.3456
10.1847 | -93.3452
10.185 | -93.3447
10.1853 | -93.3447
10.1856 | -93.3452
10.186 | -93.3456
10.1863 | -93.3456
10.1866 | -93.3455
10.1869 | -93.3452
10.1872 | -93.3449
10.1875 | -93.3442
10.1878 | -93.3441
10.1881 | -93.3441
10.1884 | -93.3443
10.1887 | -93.3446
10.1891 | -93.3454
10.1894 | -93.3464
10.1897 | -93.3467
10.19 | -93.3466
10.1903 | -93.3463
10.1906 | -93.346
10.1909 | -93.3462
10.1912 | -93.3457
10.1915 | -93.3455
10.1919 | -93.3452
10.1922 | -93.3446
10.1925 | -93.3443
10.1928 | -93.3444
10.1931 | -93.3448
10.1934 | -93.3446
10.1937 | -93.3444
10.194 | -93.3443
10.1943 | -93.3444
10.1946 | -93.3445
10.195 | -93.3447
10.1953 | -93.3446
10.1956 | -93.3454
10.1959 | -93.3457
10.1962 | -93.3456
10.1965 | -93.3454
10.1968 | -93.3452
10.1971 | -93.3453
10.1974 | -93.3453
10.1977 | -93.3451
10.1981 | -93.3447
10.1984 | -93.3448
10.1987 | -93.3449
10.199 | -93.3453
10.1993 | -93.3452
10.1996 | -93.3444
10.1999 | -93.3439
10.2002 | -93.3437
10.2005 | -93.3439
10.2009 | -93.3437
10.2012 | -93.3431
10.2015 | -93.3427
10.2018 | -93.3425
10.2021 | -93.3423
10.2024 | -93.3421
10.2027 | -93.3419
10.203 | -93.3417
10.2033 | -93.3416
10.2037 | -93.3415
10.204 | -93.3414
10.2043 | -93.3414
10.2046 | -93.3412
10.2049 | -93.3411
10.2052 | -93.3408
10.2055 | -93.3404
10.2058 | -93.3403
10.2061 | -93.3406
10.2064 | -93.3407
10.2068 | -93.3408
10.2071 | -93.341
10.2074 | -93.3417
10.2077 | -93.3417
10.208 | -93.3415
10.2083 | -93.3414
10.2086 | -93.3417
10.2089 | -93.3418
10.2092 | -93.3422
10.2096 | -93.3424
10.2099 | -93.3419
10.2102 | -93.3417
10.2105 | -93.3417
10.2108 | -93.3418
10.2111 | -93.3416
10.2114 | -93.3414
10.2117 | -93.3414
10.212 | -93.3416
10.2123 | -93.3415
10.2127 | -93.3411
10.213 | -93.3408
10.2133 | -93.3407
10.2136 | -93.3408
10.2139 | -93.341
10.2142 | -93.3404
10.2145 | -93.3402
10.2148 | -93.3401
10.2151 | -93.3399
10.2155 | -93.3393
10.2158 | -93.339
10.2161 | -93.3389
10.2164 | -93.3386
10.2167 | -93.3383
10.217 | -93.3383
10.2173 | -93.3385
10.2176 | -93.3385
10.2179 | -93.3382
10.2182 | -93.3382
10.2186 | -93.3386
10.2189 | -93.3389
10.2192 | -93.3391
10.2195 | -93.3394
10.2198 | -93.3395
10.2201 | -93.3388
10.2204 | -93.3383
10.2207 | -93.3383
10.221 | -93.3383
10.2214 | -93.3385
10.2217 | -93.3388
10.222 | -93.3387
10.2223 | -93.3383
10.2226 | -93.338
10.2229 | -93.3383
10.2232 | -93.3385
10.2235 | -93.3387
10.2238 | -93.3386
10.2241 | -93.3385
10.2245 | -93.3384
10.2248 | -93.3383
10.2251 | -93.3381
10.2254 | -93.338
10.2257 | -93.3377
10.226 | -93.337
10.2263 | -93.3369
10.2266 | -93.3372
10.2269 | -93.3376
10.2273 | -93.3374
10.2276 | -93.3371
10.2279 | -93.3366
10.2282 | -93.3359
10.2285 | -93.3358
10.2288 | -93.3356
10.2291 | -93.3358
10.2294 | -93.3359
10.2297 | -93.3358
10.23 | -93.336
10.2304 | -93.3366
10.2307 | -93.3368
10.231 | -93.3368
10.2313 | -93.3375
10.2316 | -93.3383
10.2319 | -93.3385
10.2322 | -93.3386
10.2325 | -93.3385
10.2328 | -93.3385
10.2332 | -93.3385
10.2335 | -93.3386
10.2338 | -93.3388
10.2341 | -93.3394
10.2344 | -93.3392
10.2347 | -93.3388
10.235 | -93.3389
10.2353 | -93.3391
10.2356 | -93.3387
10.2359 | -93.3385
10.2363 | -93.3385
10.2366 | -93.3387
10.2369 | -93.3391
10.2372 | -93.3392
10.2375 | -93.339
10.2378 | -93.3389
10.2381 | -93.3389
10.2384 | -93.3388
10.2387 | -93.3389
10.2391 | -93.3392
10.2394 | -93.3393
10.2397 | -93.3392
10.24 | -93.3391
10.2403 | -93.3391
10.2406 | -93.339
10.2409 | -93.3392
10.2412 | -93.3393
10.2415 | -93.3393
10.2418 | -93.3392
10.2422 | -93.3389
10.2425 | -93.339
10.2428 | -93.3393
10.2431 | -93.3393
10.2434 | -93.339
10.2437 | -93.3391
10.244 | -93.339
10.2443 | -93.3392
10.2446 | -93.3392
10.245 | -93.3389
10.2453 | -93.339
10.2456 | -93.339
10.2459 | -93.3389
10.2462 | -93.339
10.2465 | -93.3396
10.2468 | -93.34
10.2471 | -93.34
10.2474 | -93.3399
10.2477 | -93.3399
10.2481 | -93.3398
10.2484 | -93.3396
10.2487 | -93.3397
10.249 | -93.3397
10.2493 | -93.3394
10.2496 | -93.3394
10.2499 | -93.339
10.2502 | -93.3389
10.2505 | -93.3389
10.2509 | -93.3388
10.2512 | -93.3386
10.2515 | -93.3384
10.2518 | -93.3381
10.2521 | -93.3381
10.2524 | -93.3384
10.2527 | -93.3383
10.253 | -93.338
10.2533 | -93.3379
10.2536 | -93.3378
10.254 | -93.3378
10.2543 | -93.3384
10.2546 | -93.3388
10.2549 | -93.3388
10.2552 | -93.3387
10.2555 | -93.3388
10.2558 | -93.3385
10.2561 | -93.3383
10.2564 | -93.3379
10.2568 | -93.3379
10.2571 | -93.3377
10.2574 | -93.3374
10.2577 | -93.3379
10.258 | -93.3387
10.2583 | -93.3391
10.2586 | -93.3395
10.2589 | -93.3396
10.2592 | -93.3401
10.2595 | -93.3403
10.2599 | -93.3402
10.2602 | -93.3402
10.2605 | -93.34
10.2608 | -93.3407
10.2611 | -93.3415
10.2614 | -93.3412
10.2617 | -93.3414
10.262 | -93.3415
10.2623 | -93.3419
10.2627 | -93.342
10.263 | -93.3419
10.2633 | -93.3421
10.2636 | -93.3416
10.2639 | -93.3405
10.2642 | -93.3399
10.2645 | -93.3398
10.2648 | -93.34
10.2651 | -93.3397
10.2654 | -93.3396
10.2658 | -93.3397
10.2661 | -93.3398
10.2664 | -93.3399
10.2667 | -93.3404
10.267 | -93.3408
10.2673 | -93.341
10.2676 | -93.3413
10.2679 | -93.3412
10.2682 | -93.3413
10.2686 | -93.3417
10.2689 | -93.3421
10.2692 | -93.3421
10.2695 | -93.342
10.2698 | -93.3419
10.2701 | -93.3418
10.2704 | -93.3418
10.2707 | -93.342
10.271 | -93.3418
10.2713 | -93.3418
10.2717 | -93.3423
10.272 | -93.3426
10.2723 | -93.3426
10.2726 | -93.3428
10.2729 | -93.3427
10.2732 | -93.3427
10.2735 | -93.3428
10.2738 | -93.3428
10.2741 | -93.343
10.2745 | -93.3429
10.2748 | -93.3427
10.2751 | -93.3424
10.2754 | -93.3425
10.2757 | -93.3424
10.276 | -93.3424
10.2763 | -93.3425
10.2766 | -93.343
10.2769 | -93.3431
10.2772 | -93.3435
10.2776 | -93.344
10.2779 | -93.344
10.2782 | -93.3439
10.2785 | -93.3434
10.2788 | -93.3432
10.2791 | -93.3432
10.2794 | -93.3435
10.2797 | -93.3434
10.28 | -93.3431
10.2804 | -93.343
10.2807 | -93.3427
10.281 | -93.3427
10.2813 | -93.3429
10.2816 | -93.3429
10.2819 | -93.3431
10.2822 | -93.3435
10.2825 | -93.3433
10.2828 | -93.3433
10.2831 | -93.3434
10.2835 | -93.3432
10.2838 | -93.3433
10.2841 | -93.3432
10.2844 | -93.343
10.2847 | -93.3433
10.285 | -93.344
10.2853 | -93.3448
10.2856 | -93.3453
10.2859 | -93.3454
10.2863 | -93.3454
10.2866 | -93.3452
10.2869 | -93.3451
10.2872 | -93.3451
10.2875 | -93.345
10.2878 | -93.345
10.2881 | -93.345
10.2884 | -93.3448
10.2887 | -93.3448
10.289 | -93.3452
10.2894 | -93.3451
10.2897 | -93.345
10.29 | -93.3451
10.2903 | -93.3451
10.2906 | -93.3455
10.2909 | -93.3456
10.2912 | -93.3455
10.2915 | -93.3456
10.2918 | -93.3458
10.2922 | -93.3459
10.2925 | -93.3461
10.2928 | -93.346
10.2931 | -93.3462
10.2934 | -93.3462
10.2937 | -93.3462
10.294 | -93.3461
10.2943 | -93.3462
10.2946 | -93.3458
10.2949 | -93.3455
10.2953 | -93.345
10.2956 | -93.3448
10.2959 | -93.3449
10.2962 | -93.3449
10.2965 | -93.3454
10.2968 | -93.346
10.2971 | -93.3465
10.2974 | -93.3466
10.2977 | -93.3463
10.2981 | -93.346
10.2984 | -93.3461
10.2987 | -93.3463
10.299 | -93.3464
10.2993 | -93.3469
10.2996 | -93.3472
10.2999 | -93.3476
10.3002 | -93.3479
10.3005 | -93.3485
10.3008 | -93.3493
10.3012 | -93.3497
10.3015 | -93.3499
10.3018 | -93.3503
10.3021 | -93.3504
10.3024 | -93.3502
10.3027 | -93.3509
10.303 | -93.3514
10.3033 | -93.3517
10.3036 | -93.3515
10.304 | -93.352
10.3043 | -93.3521
10.3046 | -93.352
10.3049 | -93.3518
10.3052 | -93.3523
10.3055 | -93.3526
10.3058 | -93.3528
10.3061 | -93.3529
10.3064 | -93.353
10.3067 | -93.3535
10.3071 | -93.3535
10.3074 | -93.3535
10.3077 | -93.3531
10.308 | -93.3526
10.3083 | -93.3528
10.3086 | -93.353
10.3089 | -93.3534
10.3092 | -93.3536
10.3095 | -93.3533
10.3099 | -93.3535
10.3102 | -93.3541
10.3105 | -93.3549
10.3108 | -93.3553
10.3111 | -93.3554
10.3114 | -93.3556
10.3117 | -93.3558
10.312 | -93.3561
10.3123 | -93.3565
10.3126 | -93.3569
10.313 | -93.3568
10.3133 | -93.357
10.3136 | -93.3573
10.3139 | -93.3575
10.3142 | -93.3576
10.3145 | -93.358
10.3148 | -93.358
10.3151 | -93.3582
10.3154 | -93.3589
10.3158 | -93.3592
10.3161 | -93.3593
10.3164 | -93.3593
10.3167 | -93.3595
10.317 | -93.3594
10.3173 | -93.3593
10.3176 | -93.3593
10.3179 | -93.3595
10.3182 | -93.3597
10.3186 | -93.36
10.3189 | -93.3607
10.3192 | -93.3613
10.3195 | -93.3615
10.3198 | -93.3617
10.3201 | -93.3618
10.3204 | -93.3619
10.3207 | -93.362
10.321 | -93.3621
10.3213 | -93.3623
10.3217 | -93.3626
10.322 | -93.3628
10.3223 | -93.3632
10.3226 | -93.3632
10.3229 | -93.363
10.3232 | -93.3629
10.3235 | -93.3629
10.3238 | -93.3633
10.3241 | -93.3642
10.3245 | -93.3645
10.3248 | -93.3645
10.3251 | -93.3647
10.3254 | -93.3651
10.3257 | -93.3648
10.326 | -93.365
10.3263 | -93.3652
10.3266 | -93.3653
10.3269 | -93.3663
10.3272 | -93.3666
10.3276 | -93.3668
10.3279 | -93.3674
10.3282 | -93.3677
10.3285 | -93.3681
10.3288 | -93.3683
10.3291 | -93.3686
10.3294 | -93.3687
10.3297 | -93.3682
10.33 | -93.3682
10.3303 | -93.3683
10.3307 | -93.3683
10.331 | -93.3687
10.3313 | -93.3691
10.3316 | -93.3693
10.3319 | -93.3695
10.3322 | -93.3698
10.3325 | -93.3705
10.3328 | -93.3709
10.3331 | -93.3711
10.3335 | -93.3715
10.3338 | -93.3718
10.3341 | -93.3723
10.3344 | -93.3729
10.3347 | -93.3732
10.335 | -93.3734
10.3353 | -93.3738
10.3356 | -93.3742
10.3359 | -93.3743
10.3362 | -93.3745
10.3366 | -93.3747
10.3369 | -93.3751
10.3372 | -93.375
10.3375 | -93.3752
10.3378 | -93.3752
10.3381 | -93.3749
10.3384 | -93.3753
10.3387 | -93.3763
10.339 | -93.3778
10.3394 | -93.3787
10.3397 | -93.3801
10.34 | -93.381
10.3403 | -93.3815
10.3406 | -93.3818
10.3409 | -93.3819
10.3412 | -93.3818
10.3415 | -93.3817
10.3418 | -93.3818
10.3422 | -93.3823
10.3425 | -93.3825
10.3428 | -93.3826
10.3431 | -93.3827
10.3434 | -93.3829
10.3437 | -93.3834
10.344 | -93.3836
10.3443 | -93.3838
10.3446 | -93.3839
10.3449 | -93.3843
10.3453 | -93.385
10.3456 | -93.3854
10.3459 | -93.3858
10.3462 | -93.3861
10.3465 | -93.3862
10.3468 | -93.3865
10.3471 | -93.3866
10.3474 | -93.3863
10.3477 | -93.3862
10.3481 | -93.3879
10.3484 | -93.3889
10.3487 | -93.3898
10.349 | -93.3901
10.3493 | -93.3904
10.3496 | -93.3907
10.3499 | -93.3909
10.3502 | -93.3911
10.3505 | -93.3913
10.3508 | -93.3912
10.3512 | -93.3911
10.3515 | -93.3917
10.3518 | -93.3923
10.3521 | -93.3936
10.3524 | -93.3939
10.3527 | -93.3941
10.353 | -93.3944
10.3533 | -93.3949
10.3536 | -93.3956
10.354 | -93.396
10.3543 | -93.3959
10.3546 | -93.3956
10.3549 | -93.3957
10.3552 | -93.3961
10.3555 | -93.3966
10.3558 | -93.3971
10.3561 | -93.3973
10.3564 | -93.3979
10.3567 | -93.398
10.3571 | -93.3984
10.3574 | -93.3991
10.3577 | -93.3991
10.358 | -93.3997
10.3583 | -93.4001
10.3586 | -93.4002
10.3589 | -93.3997
10.3592 | -93.3989
10.3595 | -93.3991
10.3599 | -93.3994
10.3602 | -93.3999
10.3605 | -93.4001
10.3608 | -93.4004
10.3611 | -93.4008
10.3614 | -93.4013
10.3617 | -93.4015
10.362 | -93.4017
10.3623 | -93.4024
10.3626 | -93.403
10.363 | -93.4033
10.3633 | -93.4036
10.3636 | -93.4044
10.3639 | -93.4047
10.3642 | -93.4051
10.3645 | -93.4054
10.3648 | -93.4057
10.3651 | -93.4061
10.3654 | -93.4067
10.3658 | -93.4069
10.3661 | -93.4072
10.3664 | -93.408
10.3667 | -93.4086
10.367 | -93.4089
10.3673 | -93.4096
10.3676 | -93.4103
10.3679 | -93.4106
10.3682 | -93.411
10.3685 | -93.4114
10.3689 | -93.4117
10.3692 | -93.4117
10.3695 | -93.4119
10.3698 | -93.4124
10.3701 | -93.4129
10.3704 | -93.4129
10.3707 | -93.4124
10.371 | -93.4129
10.3713 | -93.4133
10.3717 | -93.4137
10.372 | -93.4142
10.3723 | -93.4149
10.3726 | -93.4155
10.3729 | -93.4158
10.3732 | -93.4162
10.3735 | -93.4166
10.3738 | -93.4169
10.3741 | -93.4171
10.3744 | -93.4175
10.3748 | -93.4178
10.3751 | -93.4181
10.3754 | -93.4191
10.3757 | -93.4193
10.376 | -93.419
10.3763 | -93.4188
10.3766 | -93.4188
10.3769 | -93.4193
10.3772 | -93.4197
10.3776 | -93.4199
10.3779 | -93.4202
10.3782 | -93.4207
10.3785 | -93.4212
10.3788 | -93.422
10.3791 | -93.4226
10.3794 | -93.4231
10.3797 | -93.4234
10.38 | -93.4236
10.3803 | -93.4241
10.3807 | -93.4242
10.381 | -93.4243
10.3813 | -93.4244
10.3816 | -93.4249
10.3819 | -93.4253
10.3822 | -93.4257
10.3825 | -93.4257
10.3828 | -93.4263
10.3831 | -93.4268
10.3835 | -93.4272
10.3838 | -93.4276
10.3841 | -93.428
10.3844 | -93.4285
10.3847 | -93.4291
10.385 | -93.4296
10.3853 | -93.4301
10.3856 | -93.4307
10.3859 | -93.4312
10.3862 | -93.4319
10.3866 | -93.4322
10.3869 | -93.4327
10.3872 | -93.4331
10.3875 | -93.4335
10.3878 | -93.434
10.3881 | -93.4341
10.3884 | -93.4338
10.3887 | -93.4338
10.389 | -93.4339
10.3894 | -93.4346
10.3897 | -93.4345
10.39 | -93.4343
10.3903 | -93.4346
10.3906 | -93.4348
10.3909 | -93.4353
10.3912 | -93.4355
10.3915 | -93.436
10.3918 | -93.4366
10.3921 | -93.4371
10.3925 | -93.4375
10.3928 | -93.4379
10.3931 | -93.439
10.3934 | -93.4396
10.3937 | -93.44
10.394 | -93.4405
10.3943 | -93.441
10.3946 | -93.4418
10.3949 | -93.442
10.3953 | -93.4425
10.3956 | -93.443
10.3959 | -93.4435
10.3962 | -93.444
10.3965 | -93.4444
10.3968 | -93.4449
10.3971 | -93.4456
10.3974 | -93.4464
10.3977 | -93.4481
10.398 | -93.4493
10.3984 | -93.4498
10.3987 | -93.4502
10.399 | -93.4505
10.3993 | -93.4508
10.3996 | -93.4514
10.3999 | -93.4519
10.4002 | -93.4523
10.4005 | -93.4535
10.4008 | -93.4544
10.4012 | -93.4549
10.4015 | -93.4554
10.4018 | -93.456
10.4021 | -93.4571
10.4024 | -93.4581
10.4027 | -93.4585
10.403 | -93.4589
10.4033 | -93.4595
10.4036 | -93.4601
10.4039 | -93.4607
10.4043 | -93.4612
10.4046 | -93.4618
10.4049 | -93.4622
10.4052 | -93.4628
10.4055 | -93.4634
10.4058 | -93.4639
10.4061 | -93.4645
10.4064 | -93.4653
10.4067 | -93.4658
10.4071 | -93.4661
10.4074 | -93.4666
10.4077 | -93.4672
10.408 | -93.4678
10.4083 | -93.4687
10.4086 | -93.4693
10.4089 | -93.4698
10.4092 | -93.4704
10.4095 | -93.471
10.4098 | -93.4715
10.4102 | -93.4725
10.4105 | -93.4732
10.4108 | -93.4736
10.4111 | -93.4746
10.4114 | -93.4751
10.4117 | -93.4754
10.412 | -93.476
10.4123 | -93.4765
10.4126 | -93.4768
10.413 | -93.4768
10.4133 | -93.4773
10.4136 | -93.4778
10.4139 | -93.4784
10.4142 | -93.4794
10.4145 | -93.4802
10.4148 | -93.481
10.4151 | -93.4818
10.4154 | -93.4823
10.4157 | -93.4828
10.4161 | -93.4833
10.4164 | -93.4842
10.4167 | -93.4853
10.417 | -93.486
10.4173 | -93.4865
10.4176 | -93.4872
10.4179 | -93.488
10.4182 | -93.4884
10.4185 | -93.4886
10.4189 | -93.4889
10.4192 | -93.4893
10.4195 | -93.4899
10.4198 | -93.49
10.4201 | -93.4901
10.4204 | -93.4908
10.4207 | -93.4915
10.421 | -93.492
10.4213 | -93.4928
10.4216 | -93.4936
10.422 | -93.4941
10.4223 | -93.4948
10.4226 | -93.4962
10.4229 | -93.4975
10.4232 | -93.4985
10.4235 | -93.4992
10.4238 | -93.5
10.4241 | -93.5009
10.4244 | -93.5018
10.4248 | -93.5021
10.4251 | -93.5025
10.4254 | -93.503
10.4257 | -93.5034
10.426 | -93.5035
10.4263 | -93.5041
10.4266 | -93.5049
10.4269 | -93.5056
10.4272 | -93.5063
10.4275 | -93.5062
10.4279 | -93.5065
10.4282 | -93.5073
10.4285 | -93.5081
10.4288 | -93.5091
10.4291 | -93.51
10.4294 | -93.5109
10.4297 | -93.5119
10.43 | -93.5124
10.4303 | -93.5133
10.4307 | -93.514
10.431 | -93.5137
10.4313 | -93.5141
10.4316 | -93.5146
10.4319 | -93.5152
10.4322 | -93.5159
10.4325 | -93.5164
10.4328 | -93.5171
10.4331 | -93.5177
10.4334 | -93.5188
10.4338 | -93.5199
10.4341 | -93.5205
10.4344 | -93.5209
10.4347 | -93.5214
10.435 | -93.5219
10.4353 | -93.5225
10.4356 | -93.5233
10.4359 | -93.5242
10.4362 | -93.5248
10.4366 | -93.5252
10.4369 | -93.5256
10.4372 | -93.526
10.4375 | -93.5266
10.4378 | -93.5273
10.4381 | -93.5284
10.4384 | -93.5291
10.4387 | -93.5296
10.439 | -93.5301
10.4393 | -93.5308
10.4397 | -93.5311
10.44 | -93.5315
10.4403 | -93.5317
10.4406 | -93.5319
10.4409 | -93.5325
10.4412 | -93.5329
10.4415 | -93.5335
10.4418 | -93.5341
10.4421 | -93.5351
10.4425 | -93.5356
10.4428 | -93.5359
10.4431 | -93.5366
10.4434 | -93.5373
10.4437 | -93.5378
10.444 | -93.5387
10.4443 | -93.5393
10.4446 | -93.5399
10.4449 | -93.5411
10.4452 | -93.5419
10.4456 | -93.5426
10.4459 | -93.5433
10.4462 | -93.544
10.4465 | -93.5448
10.4468 | -93.5452
10.4471 | -93.5458
10.4474 | -93.5464
10.4477 | -93.5471
10.448 | -93.5478
10.4484 | -93.5484
10.4487 | -93.5491
10.449 | -93.5499
10.4493 | -93.5504
10.4496 | -93.551
10.4499 | -93.5515
10.4502 | -93.5524
10.4505 | -93.5532
10.4508 | -93.5541
10.4511 | -93.5551
10.4515 | -93.5559
10.4518 | -93.5573
10.4521 | -93.5588
10.4524 | -93.5599
10.4527 | -93.5613
10.453 | -93.562
10.4533 | -93.5625
10.4536 | -93.5638
10.4539 | -93.5644
10.4543 | -93.5651
10.4546 | -93.5656
10.4549 | -93.5661
10.4552 | -93.5669
10.4555 | -93.5675
10.4558 | -93.5683
10.4561 | -93.5691
10.4564 | -93.5699
10.4567 | -93.5707
10.4571 | -93.5717
10.4574 | -93.5727
10.4577 | -93.5738
10.458 | -93.5748
10.4583 | -93.5754
10.4586 | -93.5757
10.4589 | -93.5764
10.4592 | -93.577
10.4595 | -93.5776
10.4598 | -93.5782
10.4602 | -93.5789
10.4605 | -93.5796
10.4608 | -93.5802
10.4611 | -93.5812
10.4614 | -93.5816
10.4617 | -93.5824
10.462 | -93.583
10.4623 | -93.5837
10.4626 | -93.584
10.4629 | -93.5843
10.4633 | -93.5851
10.4636 | -93.5856
10.4639 | -93.5862
10.4642 | -93.5872
10.4645 | -93.588
10.4648 | -93.5891
10.4651 | -93.5897
10.4654 | -93.5904
10.4657 | -93.5911
10.4661 | -93.5917
10.4664 | -93.5917
10.4667 | -93.5917
10.467 | -93.5926
10.4673 | -93.5943
10.4676 | -93.5954
10.4679 | -93.596
10.4682 | -93.5966
10.4685 | -93.5975
10.4688 | -93.5984
10.4692 | -93.5997
10.4695 | -93.6006
10.4698 | -93.6013
10.4701 | -93.6021
10.4704 | -93.6034
10.4707 | -93.6047
10.471 | -93.6054
10.4713 | -93.606
10.4716 | -93.6065
10.472 | -93.6072
10.4723 | -93.6078
10.4726 | -93.6083
10.4729 | -93.609
10.4732 | -93.6098
10.4735 | -93.611
10.4738 | -93.6119
10.4741 | -93.6126
10.4744 | -93.6132
10.4748 | -93.6141
10.4751 | -93.6149
10.4754 | -93.6154
10.4757 | -93.6163
10.476 | -93.6176
10.4763 | -93.6198
10.4766 | -93.6215
10.4769 | -93.6226
10.4772 | -93.6232
10.4775 | -93.6235
10.4779 | -93.6236
10.4782 | -93.6242
10.4785 | -93.625
10.4788 | -93.6257
10.4791 | -93.6262
10.4794 | -93.6269
10.4797 | -93.6277
10.48 | -93.6283
10.4803 | -93.6288
10.4807 | -93.6293
10.481 | -93.6298
10.4813 | -93.6304
10.4816 | -93.631
10.4819 | -93.6316
10.4822 | -93.632
10.4825 | -93.6326
10.4828 | -93.6335
10.4831 | -93.6343
10.4834 | -93.635
10.4838 | -93.6357
10.4841 | -93.6365
10.4844 | -93.6372
10.4847 | -93.6377
10.485 | -93.6383
10.4853 | -93.6389
10.4856 | -93.6393
10.4859 | -93.6398
10.4862 | -93.6405
10.4866 | -93.6411
10.4869 | -93.6419
10.4872 | -93.6425
10.4875 | -93.6432
10.4878 | -93.6437
10.4881 | -93.6441
10.4884 | -93.6453
10.4887 | -93.6467
10.489 | -93.6477
10.4893 | -93.6483
10.4897 | -93.6492
10.49 | -93.65
10.4903 | -93.6508
10.4906 | -93.6514
10.4909 | -93.6517
10.4912 | -93.6527
10.4915 | -93.6542
10.4918 | -93.6552
10.4921 | -93.6561
10.4925 | -93.657
10.4928 | -93.658
10.4931 | -93.6587
10.4934 | -93.6592
10.4937 | -93.6592
10.494 | -93.6597
10.4943 | -93.6607
10.4946 | -93.6615
10.4949 | -93.6625
10.4952 | -93.6633
10.4956 | -93.664
10.4959 | -93.6651
10.4962 | -93.6662
10.4965 | -93.667
10.4968 | -93.6677
10.4971 | -93.6684
10.4974 | -93.6685
10.4977 | -93.6689
10.498 | -93.6696
10.4984 | -93.6711
10.4987 | -93.6719
10.499 | -93.6724
10.4993 | -93.6729
10.4996 | -93.6735
10.4999 | -93.6741
10.5002 | -93.6749
10.5005 | -93.6757
10.5008 | -93.6766
10.5011 | -93.6773
10.5015 | -93.678
10.5018 | -93.6787
10.5021 | -93.6794
10.5024 | -93.6799
10.5027 | -93.68
10.503 | -93.6802
10.5033 | -93.6807
10.5036 | -93.6811
10.5039 | -93.6818
10.5043 | -93.6825
10.5046 | -93.6831
10.5049 | -93.6838
10.5052 | -93.6847
10.5055 | -93.6858
10.5058 | -93.6872
10.5061 | -93.688
10.5064 | -93.6887
10.5067 | -93.6894
10.507 | -93.6901
10.5074 | -93.6907
10.5077 | -93.6913
10.508 | -93.6924
10.5083 | -93.6942
10.5086 | -93.696
10.5089 | -93.6969
10.5092 | -93.6978
10.5095 | -93.6984
10.5098 | -93.6989
10.5102 | -93.6995
10.5105 | -93.6997
10.5108 | -93.7001
10.5111 | -93.7008
10.5114 | -93.7015
10.5117 | -93.7022
10.512 | -93.7028
10.5123 | -93.7035
10.5126 | -93.7044
10.5129 | -93.7051
10.5133 | -93.7058
10.5136 | -93.7061
10.5139 | -93.7065
10.5142 | -93.7072
10.5145 | -93.7078
10.5148 | -93.7087
10.5151 | -93.71
10.5154 | -93.7114
10.5157 | -93.7131
10.5161 | -93.7148
10.5164 | -93.7155
10.5167 | -93.7164
10.517 | -93.7167
10.5173 | -93.7174
10.5176 | -93.718
10.5179 | -93.7186
10.5182 | -93.7192
10.5185 | -93.7202
10.5188 | -93.7211
10.5192 | -93.7219
10.5195 | -93.7225
10.5198 | -93.7234
10.5201 | -93.7243
10.5204 | -93.7254
10.5207 | -93.7263
10.521 | -93.7272
10.5213 | -93.7276
10.5216 | -93.7283
10.522 | -93.7293
10.5223 | -93.7297
10.5226 | -93.7309
10.5229 | -93.7318
10.5232 | -93.7328
10.5235 | -93.734
10.5238 | -93.7342
10.5241 | -93.7346
10.5244 | -93.7354
10.5247 | -93.7361
10.5251 | -93.7368
10.5254 | -93.7375
10.5257 | -93.7382
10.526 | -93.739
10.5263 | -93.7393
10.5266 | -93.7393
10.5269 | -93.7398
10.5272 | -93.7403
10.5275 | -93.741
10.5279 | -93.742
10.5282 | -93.7429
10.5285 | -93.7434
10.5288 | -93.7445
10.5291 | -93.7456
10.5294 | -93.7464
10.5297 | -93.747
10.53 | -93.7479
10.5303 | -93.7491
10.5306 | -93.7498
10.531 | -93.7501
10.5313 | -93.7502
10.5316 | -93.7509
10.5319 | -93.7516
10.5322 | -93.7523
10.5325 | -93.7531
10.5328 | -93.7537
10.5331 | -93.7547
10.5334 | -93.7554
10.5338 | -93.7563
10.5341 | -93.7572
10.5344 | -93.758
10.5347 | -93.7587
10.535 | -93.7593
10.5353 | -93.7599
10.5356 | -93.7606
10.5359 | -93.7616
10.5362 | -93.7625
10.5365 | -93.7632
10.5369 | -93.7641
10.5372 | -93.7651
10.5375 | -93.766
10.5378 | -93.767
10.5381 | -93.7679
10.5384 | -93.7689
10.5387 | -93.77
10.539 | -93.7707
10.5393 | -93.7718
10.5397 | -93.7727
10.54 | -93.7729
10.5403 | -93.7731
10.5406 | -93.7741
10.5409 | -93.7747
10.5412 | -93.7751
10.5415 | -93.7756
10.5418 | -93.7763
10.5421 | -93.7768
10.5424 | -93.7771
10.5428 | -93.7776
10.5431 | -93.778
10.5434 | -93.778
10.5437 | -93.7784
10.544 | -93.7793
10.5443 | -93.7799
10.5446 | -93.7806
10.5449 | -93.7813
10.5452 | -93.7821
10.5456 | -93.7831
10.5459 | -93.784
10.5462 | -93.7845
10.5465 | -93.785
10.5468 | -93.7854
10.5471 | -93.7858
10.5474 | -93.7865
10.5477 | -93.7874
10.548 | -93.7882
10.5483 | -93.7891
10.5487 | -93.7898
10.549 | -93.7905
10.5493 | -93.7911
10.5496 | -93.7915
10.5499 | -93.7918
10.5502 | -93.7921
10.5505 | -93.7928
10.5508 | -93.7931
10.5511 | -93.7936
10.5515 | -93.7941
10.5518 | -93.7947
10.5521 | -93.7955
10.5524 | -93.7961
10.5527 | -93.7967
10.553 | -93.797
10.5533 | -93.7972
10.5536 | -93.7979
10.5539 | -93.7987
10.5542 | -93.7995
10.5546 | -93.7998
10.5549 | -93.801
10.5552 | -93.8019
10.5555 | -93.8025
10.5558 | -93.8032
10.5561 | -93.804
10.5564 | -93.8045
10.5567 | -93.8047
10.557 | -93.8054
10.5574 | -93.8064
10.5577 | -93.8071
10.558 | -93.8077
10.5583 | -93.8083
10.5586 | -93.8089
10.5589 | -93.8097
10.5592 | -93.8103
10.5595 | -93.8106
10.5598 | -93.8112
10.5601 | -93.8118
10.5605 | -93.8125
10.5608 | -93.8132
10.5611 | -93.8139
10.5614 | -93.8144
10.5617 | -93.8147
10.562 | -93.8152
10.5623 | -93.8158
10.5626 | -93.8167
10.5629 | -93.8173
10.5633 | -93.8176
10.5636 | -93.8184
10.5639 | -93.8193
10.5642 | -93.82
10.5645 | -93.8207
10.5648 | -93.8214
10.5651 | -93.8222
10.5654 | -93.8227
10.5657 | -93.8233
10.566 | -93.8239
10.5664 | -93.8247
10.5667 | -93.8253
10.567 | -93.8259
10.5673 | -93.8266
10.5676 | -93.8272
10.5679 | -93.8279
10.5682 | -93.8288
10.5685 | -93.8297
10.5688 | -93.8304
10.5692 | -93.8315
10.5695 | -93.8321
10.5698 | -93.8327
10.5701 | -93.8332
10.5704 | -93.8339
10.5707 | -93.8343
10.571 | -93.8349
10.5713 | -93.836
10.5716 | -93.8367
10.5719 | -93.8375
10.5723 | -93.8384
10.5726 | -93.8386
10.5729 | -93.839
10.5732 | -93.8395
10.5735 | -93.8403
10.5738 | -93.8411
10.5741 | -93.8423
10.5744 | -93.8436
10.5747 | -93.8443
10.5751 | -93.8452
10.5754 | -93.8469
10.5757 | -93.8478
10.576 | -93.8483
10.5763 | -93.849
10.5766 | -93.8498
10.5769 | -93.8505
10.5772 | -93.8511
10.5775 | -93.8517
10.5778 | -93.8524
10.5782 | -93.8529
10.5785 | -93.8539
10.5788 | -93.8551
10.5791 | -93.8563
10.5794 | -93.8571
10.5797 | -93.8578
10.58 | -93.8588
10.5803 | -93.8595
10.5806 | -93.8603
10.581 | -93.8613
10.5813 | -93.8614
10.5816 | -93.8617
10.5819 | -93.8622
10.5822 | -93.863
10.5825 | -93.8639
10.5828 | -93.8652
10.5831 | -93.8662
10.5834 | -93.867
10.5837 | -93.8676
10.5841 | -93.8683
10.5844 | -93.8689
10.5847 | -93.8697
10.585 | -93.8705
10.5853 | -93.8709
10.5856 | -93.872
10.5859 | -93.8727
10.5862 | -93.8734
10.5865 | -93.8736
10.5869 | -93.8739
10.5872 | -93.8745
10.5875 | -93.8756
10.5878 | -93.876
10.5881 | -93.876
10.5884 | -93.8764
10.5887 | -93.8776
10.589 | -93.8784
10.5893 | -93.879
10.5896 | -93.8796
10.59 | -93.88
10.5903 | -93.8804
10.5906 | -93.8812
10.5909 | -93.882
10.5912 | -93.8827
10.5915 | -93.8833
10.5918 | -93.8839
10.5921 | -93.8846
10.5924 | -93.8853
10.5928 | -93.8866
10.5931 | -93.8878
10.5934 | -93.8887
10.5937 | -93.8892
10.594 | -93.8896
10.5943 | -93.8909
10.5946 | -93.892
10.5949 | -93.8927
10.5952 | -93.8934
10.5955 | -93.8941
10.5959 | -93.8947
10.5962 | -93.8949
10.5965 | -93.8953
10.5968 | -93.8956
10.5971 | -93.8962
10.5974 | -93.8974
10.5977 | -93.8981
10.598 | -93.8985
10.5983 | -93.8987
10.5987 | -93.8991
10.599 | -93.8996
10.5993 | -93.9001
10.5996 | -93.9005
10.5999 | -93.9009
10.6002 | -93.9019
10.6005 | -93.9028
10.6008 | -93.9034
10.6011 | -93.904
10.6014 | -93.9046
10.6018 | -93.905
10.6021 | -93.9047
10.6024 | -93.9049
10.6027 | -93.9056
10.603 | -93.9066
10.6033 | -93.9075
10.6036 | -93.9087
10.6039 | -93.9099
10.6042 | -93.9106
10.6046 | -93.9111
10.6049 | -93.912
10.6052 | -93.9127
10.6055 | -93.913
10.6058 | -93.9137
10.6061 | -93.9157
10.6064 | -93.9174
10.6067 | -93.9186
10.607 | -93.9194
10.6074 | -93.9198
10.6077 | -93.9198
10.608 | -93.9205
10.6083 | -93.9212
10.6086 | -93.9221
10.6089 | -93.9228
10.6092 | -93.9234
10.6095 | -93.924
10.6098 | -93.9246
10.6101 | -93.9253
10.6105 | -93.9259
10.6108 | -93.9265
10.6111 | -93.9271
10.6114 | -93.9277
10.6117 | -93.9283
10.612 | -93.9289
10.6123 | -93.9292
10.6126 | -93.9296
10.6129 | -93.9303
10.6133 | -93.9308
10.6136 | -93.9314
10.6139 | -93.932
10.6142 | -93.9327
10.6145 | -93.9332
10.6148 | -93.9339
10.6151 | -93.9345
10.6154 | -93.9352
10.6157 | -93.9358
10.616 | -93.9371
10.6164 | -93.9382
10.6167 | -93.9388
10.617 | -93.9392
10.6173 | -93.9399
10.6176 | -93.9405
10.6179 | -93.9408
10.6182 | -93.9414
10.6185 | -93.9422
10.6188 | -93.9434
10.6191 | -93.9453
10.6195 | -93.9467
10.6198 | -93.9473
10.6201 | -93.948
10.6204 | -93.9486
10.6207 | -93.9489
10.621 | -93.9497
10.6213 | -93.9503
10.6216 | -93.9505
10.6219 | -93.95
10.6223 | -93.9498
10.6226 | -93.95
10.6229 | -93.9503
10.6232 | -93.9508
10.6235 | -93.9515
10.6238 | -93.9519
10.6241 | -93.9518
10.6244 | -93.9523
10.6247 | -93.953
10.625 | -93.9534
10.6254 | -93.9538
10.6257 | -93.9543
10.626 | -93.9547
10.6263 | -93.9549
10.6266 | -93.9556
10.6269 | -93.9561
10.6272 | -93.9564
10.6275 | -93.9569
10.6278 | -93.9578
10.6282 | -93.9586
10.6285 | -93.9592
10.6288 | -93.9599
10.6291 | -93.9602
10.6294 | -93.9606
10.6297 | -93.9608
10.63 | -93.9611
10.6303 | -93.9618
10.6306 | -93.9623
10.631 | -93.963
10.6313 | -93.9636
10.6316 | -93.9642
10.6319 | -93.9646
10.6322 | -93.965
10.6325 | -93.9656
10.6328 | -93.9661
10.6331 | -93.9666
10.6334 | -93.9675
10.6337 | -93.9684
10.6341 | -93.969
10.6344 | -93.9696
10.6347 | -93.9703
10.635 | -93.971
10.6353 | -93.9717
10.6356 | -93.9725
10.6359 | -93.973
10.6362 | -93.974
10.6365 | -93.9751
10.6369 | -93.9756
10.6372 | -93.976
10.6375 | -93.9765
10.6378 | -93.9769
10.6381 | -93.9774
10.6384 | -93.9776
10.6387 | -93.9781
10.639 | -93.9786
10.6393 | -93.9795
10.6396 | -93.9811
10.64 | -93.9813
10.6403 | -93.9817
10.6406 | -93.9822
10.6409 | -93.9825
10.6412 | -93.9833
10.6415 | -93.9841
10.6418 | -93.9851
10.6421 | -93.986
10.6424 | -93.9866
10.6428 | -93.9872
10.6431 | -93.988
10.6434 | -93.9896
10.6437 | -93.9905
10.644 | -93.9906
10.6443 | -93.9911
10.6446 | -93.9917
10.6449 | -93.9922
10.6452 | -93.9926
10.6455 | -93.993
10.6459 | -93.9935
10.6462 | -93.9942
10.6465 | -93.9948
10.6468 | -93.9952
10.6471 | -93.9959
10.6474 | -93.9962
10.6477 | -93.9968
10.648 | -93.9977
10.6483 | -93.9981
10.6487 | -93.9984
10.649 | -93.9991
10.6493 | -94
10.6496 | -94.0007
10.6499 | -94.0017
10.6502 | -94.002
10.6505 | -94.0028
10.6508 | -94.0035
10.6511 | -94.0042
10.6514 | -94.0054
10.6518 | -94.006
10.6521 | -94.006
10.6524 | -94.0063
10.6527 | -94.0069
10.653 | -94.0071
10.6533 | -94.0076
10.6536 | -94.0081
10.6539 | -94.0085
10.6542 | -94.0082
10.6546 | -94.008
10.6549 | -94.0085
10.6552 | -94.009
10.6555 | -94.0095
10.6558 | -94.0103
10.6561 | -94.011
10.6564 | -94.0107
10.6567 | -94.01
10.657 | -94.0096
10.6573 | -94.01
10.6577 | -94.0094
10.658 | -94.0093
10.6583 | -94.0094
10.6586 | -94.0091
10.6589 | -94.0093
10.6592 | -94.01
10.6595 | -94.0105
10.6598 | -94.0111
10.6601 | -94.0119
10.6605 | -94.0122
10.6608 | -94.0129
10.6611 | -94.0143
10.6614 | -94.0156
10.6617 | -94.016
10.662 | -94.0165
10.6623 | -94.0172
10.6626 | -94.0178
10.6629 | -94.0182
10.6632 | -94.0185
10.6636 | -94.0192
10.6639 | -94.0196
10.6642 | -94.02
10.6645 | -94.0207
10.6648 | -94.0214
10.6651 | -94.0214
10.6654 | -94.0215
10.6657 | -94.0221
10.666 | -94.0222
10.6664 | -94.0227
10.6667 | -94.023
10.667 | -94.0233
10.6673 | -94.0236
10.6676 | -94.024
10.6679 | -94.0248
10.6682 | -94.0257
10.6685 | -94.0265
10.6688 | -94.0268
10.6691 | -94.0275
10.6695 | -94.0281
10.6698 | -94.0284
10.6701 | -94.0288
10.6704 | -94.0291
10.6707 | -94.0294
10.671 | -94.029
10.6713 | -94.0293
10.6716 | -94.0297
10.6719 | -94.0302
10.6723 | -94.0309
10.6726 | -94.031
10.6729 | -94.031
10.6732 | -94.0304
10.6735 | -94.03
10.6738 | -94.0303
10.6741 | -94.0307
10.6744 | -94.0309
10.6747 | -94.0313
10.675 | -94.0314
10.6754 | -94.0318
10.6757 | -94.032
10.676 | -94.0324
10.6763 | -94.0328
10.6766 | -94.0331
10.6769 | -94.0335
10.6772 | -94.0337
10.6775 | -94.0338
10.6778 | -94.0341
10.6782 | -94.0347
10.6785 | -94.0353
10.6788 | -94.0357
10.6791 | -94.0364
10.6794 | -94.0366
10.6797 | -94.037
10.68 | -94.0376
10.6803 | -94.0381
10.6806 | -94.0383
10.6809 | -94.0385
10.6813 | -94.0389
10.6816 | -94.0391
10.6819 | -94.0401
10.6822 | -94.0406
10.6825 | -94.0409
10.6828 | -94.0416
10.6831 | -94.0418
10.6834 | -94.0418
10.6837 | -94.042
10.6841 | -94.0425
10.6844 | -94.0429
10.6847 | -94.0432
10.685 | -94.0438
10.6853 | -94.0454
10.6856 | -94.047
10.6859 | -94.0473
10.6862 | -94.0478
10.6865 | -94.0483
10.6868 | -94.0486
10.6872 | -94.0487
10.6875 | -94.0493
10.6878 | -94.0493
10.6881 | -94.0489
10.6884 | -94.0489
10.6887 | -94.0488
10.689 | -94.0491
10.6893 | -94.0495
10.6896 | -94.0498
10.69 | -94.05
10.6903 | -94.0503
10.6906 | -94.0505
10.6909 | -94.0504
10.6912 | -94.0507
10.6915 | -94.0515
10.6918 | -94.0521
10.6921 | -94.0524
10.6924 | -94.0532
10.6927 | -94.0534
10.6931 | -94.0536
10.6934 | -94.0533
10.6937 | -94.0527
10.694 | -94.0529
10.6943 | -94.0541
10.6946 | -94.0546
10.6949 | -94.055
10.6952 | -94.0552
10.6955 | -94.0556
10.6959 | -94.0562
10.6962 | -94.0571
10.6965 | -94.0583
10.6968 | -94.0594
10.6971 | -94.0599
10.6974 | -94.0601
10.6977 | -94.0605
10.698 | -94.0611
10.6983 | -94.0617
10.6986 | -94.0617
10.699 | -94.0618
10.6993 | -94.0621
10.6996 | -94.0624
10.6999 | -94.062
10.7002 | -94.0619
10.7005 | -94.0622
10.7008 | -94.0623
10.7011 | -94.0627
10.7014 | -94.0629
10.7018 | -94.0634
10.7021 | -94.064
10.7024 | -94.0646
10.7027 | -94.0651
10.703 | -94.0656
10.7033 | -94.0658
10.7036 | -94.0662
10.7039 | -94.0665
10.7042 | -94.0667
10.7045 | -94.067
10.7049 | -94.0673
10.7052 | -94.0681
10.7055 | -94.0694
10.7058 | -94.0702
10.7061 | -94.0701
10.7064 | -94.0701
10.7067 | -94.0693
10.707 | -94.0689
10.7073 | -94.069
10.7077 | -94.0693
10.708 | -94.0698
10.7083 | -94.0704
10.7086 | -94.0708
10.7089 | -94.0713
10.7092 | -94.0716
10.7095 | -94.0719
10.7098 | -94.0718
10.7101 | -94.0721
10.7104 | -94.0725
10.7108 | -94.0726
10.7111 | -94.0728
10.7114 | -94.0732
10.7117 | -94.0737
10.712 | -94.0738
10.7123 | -94.0743
10.7126 | -94.0746
10.7129 | -94.075
10.7132 | -94.0752
10.7136 | -94.0754
10.7139 | -94.0758
10.7142 | -94.0764
10.7145 | -94.0771
10.7148 | -94.0773
10.7151 | -94.0777
10.7154 | -94.078
10.7157 | -94.0785
10.716 | -94.0788
10.7163 | -94.079
10.7167 | -94.0793
10.717 | -94.0792
10.7173 | -94.0794
10.7176 | -94.0797
10.7179 | -94.0798
10.7182 | -94.0798
10.7185 | -94.08
10.7188 | -94.0802
10.7191 | -94.0803
10.7195 | -94.0803
10.7198 | -94.0795
10.7201 | -94.0793
10.7204 | -94.0797
10.7207 | -94.0803
10.721 | -94.0811
10.7213 | -94.0812
10.7216 | -94.0818
10.7219 | -94.0821
10.7222 | -94.0825
10.7226 | -94.0828
10.7229 | -94.0828
10.7232 | -94.0832
10.7235 | -94.0835
10.7238 | -94.0837
10.7241 | -94.084
10.7244 | -94.0844
10.7247 | -94.0848
10.725 | -94.0855
10.7254 | -94.0857
10.7257 | -94.0859
10.726 | -94.086
10.7263 | -94.0856
10.7266 | -94.0852
10.7269 | -94.0853
10.7272 | -94.0855
10.7275 | -94.0858
10.7278 | -94.0863
10.7281 | -94.087
10.7285 | -94.0873
10.7288 | -94.0876
10.7291 | -94.0879
10.7294 | -94.088
10.7297 | -94.0881
10.73 | -94.0883
10.7303 | -94.0886
10.7306 | -94.0886
10.7309 | -94.0886
10.7313 | -94.0887
10.7316 | -94.0887
10.7319 | -94.0888
10.7322 | -94.0894
10.7325 | -94.0894
10.7328 | -94.0899
10.7331 | -94.0902
10.7334 | -94.0906
10.7337 | -94.0898
10.734 | -94.0897
10.7344 | -94.0904
10.7347 | -94.091
10.735 | -94.0915
10.7353 | -94.092
10.7356 | -94.0926
10.7359 | -94.0931
10.7362 | -94.0937
10.7365 | -94.094
10.7368 | -94.0943
10.7372 | -94.0944
10.7375 | -94.0946
10.7378 | -94.095
10.7381 | -94.0958
10.7384 | -94.0971
10.7387 | -94.0973
10.739 | -94.0971
10.7393 | -94.0974
10.7396 | -94.0979
10.7399 | -94.0975
10.7403 | -94.0969
10.7406 | -94.097
10.7409 | -94.0973
10.7412 | -94.098
10.7415 | -94.0985
10.7418 | -94.0985
10.7421 | -94.0985
10.7424 | -94.099
10.7427 | -94.0993
10.7431 | -94.0994
10.7434 | -94.0995
10.7437 | -94.0999
10.744 | -94.1001
10.7443 | -94.1004
10.7446 | -94.1008
10.7449 | -94.1015
10.7452 | -94.1018
10.7455 | -94.1019
10.7459 | -94.1019
10.7462 | -94.1028
10.7465 | -94.1036
10.7468 | -94.1039
10.7471 | -94.1038
10.7474 | -94.1034
10.7477 | -94.1035
10.748 | -94.1037
10.7483 | -94.1035
10.7486 | -94.1037
10.749 | -94.1046
10.7493 | -94.1052
10.7496 | -94.1053
10.7499 | -94.1056
10.7502 | -94.106
10.7505 | -94.1063
10.7508 | -94.1065
10.7511 | -94.1068
10.7514 | -94.1072
10.7517 | -94.1078
10.7521 | -94.108
10.7524 | -94.1079
10.7527 | -94.1085
10.753 | -94.1094
10.7533 | -94.1097
10.7536 | -94.1095
10.7539 | -94.1093
10.7542 | -94.1095
10.7545 | -94.11
10.7549 | -94.1103
10.7552 | -94.1105
10.7555 | -94.1106
10.7558 | -94.1107
10.7561 | -94.1119
10.7564 | -94.1132
10.7567 | -94.1138
10.757 | -94.1139
10.7573 | -94.114
10.7576 | -94.1142
10.758 | -94.1141
10.7583 | -94.1144
10.7586 | -94.1146
10.7589 | -94.1144
10.7592 | -94.1144
10.7595 | -94.1146
10.7598 | -94.1152
10.7601 | -94.1153
10.7604 | -94.1153
10.7608 | -94.1145
10.7611 | -94.1145
10.7614 | -94.1147
10.7617 | -94.1152
10.762 | -94.1153
10.7623 | -94.1149
10.7626 | -94.1154
10.7629 | -94.1174
10.7632 | -94.1185
10.7636 | -94.1186
10.7639 | -94.1187
10.7642 | -94.1187
10.7645 | -94.1188
10.7648 | -94.1193
10.7651 | -94.1193
10.7654 | -94.1186
10.7657 | -94.1183
10.766 | -94.1183
10.7663 | -94.1181
10.7667 | -94.1178
10.767 | -94.1179
10.7673 | -94.1184
10.7676 | -94.1185
10.7679 | -94.1185
10.7682 | -94.1182
10.7685 | -94.1184
10.7688 | -94.1192
10.7691 | -94.1197
10.7695 | -94.1198
10.7698 | -94.1195
10.7701 | -94.1195
10.7704 | -94.1196
10.7707 | -94.1197
10.771 | -94.12
10.7713 | -94.1202
10.7716 | -94.1198
10.7719 | -94.1192
10.7722 | -94.119
10.7726 | -94.1188
10.7729 | -94.1187
10.7732 | -94.119
10.7735 | -94.1197
10.7738 | -94.1204
10.7741 | -94.1207
10.7744 | -94.1205
10.7747 | -94.1205
10.775 | -94.1206
10.7754 | -94.1202
10.7757 | -94.1201
10.776 | -94.1204
10.7763 | -94.1204
10.7766 | -94.1205
10.7769 | -94.1203
10.7772 | -94.1194
10.7775 | -94.1191
10.7778 | -94.1198
10.7781 | -94.1202
10.7785 | -94.1201
10.7788 | -94.1204
10.7791 | -94.12
10.7794 | -94.1201
10.7797 | -94.1207
10.78 | -94.1214
10.7803 | -94.1213
10.7806 | -94.1213
10.7809 | -94.1206
10.7813 | -94.1195
10.7816 | -94.1195
10.7819 | -94.1195
10.7822 | -94.1195
10.7825 | -94.1195
10.7828 | -94.1197
10.7831 | -94.1195
10.7834 | -94.1195
10.7837 | -94.1195
10.784 | -94.1194
10.7844 | -94.1195
10.7847 | -94.1198
10.785 | -94.1207
10.7853 | -94.1209
10.7856 | -94.1209
10.7859 | -94.1211
10.7862 | -94.1217
10.7865 | -94.122
10.7868 | -94.1221
10.7872 | -94.1221
10.7875 | -94.1223
10.7878 | -94.1224
10.7881 | -94.1225
10.7884 | -94.1226
10.7887 | -94.1231
10.789 | -94.1237
10.7893 | -94.1237
10.7896 | -94.1239
10.7899 | -94.1239
10.7903 | -94.1239
10.7906 | -94.1241
10.7909 | -94.1246
10.7912 | -94.1246
10.7915 | -94.1249
10.7918 | -94.125
10.7921 | -94.1252
10.7924 | -94.125
10.7927 | -94.1244
10.7931 | -94.1246
10.7934 | -94.1248
10.7937 | -94.1249
10.794 | -94.125
10.7943 | -94.125
10.7946 | -94.1244
10.7949 | -94.1243
10.7952 | -94.1239
10.7955 | -94.1237
10.7958 | -94.1238
10.7962 | -94.1239
10.7965 | -94.1237
10.7968 | -94.1238
10.7971 | -94.1237
10.7974 | -94.1238
10.7977 | -94.1238
10.798 | -94.1234
10.7983 | -94.1227
10.7986 | -94.1228
10.799 | -94.1228
10.7993 | -94.1227
10.7996 | -94.1228
10.7999 | -94.1229
10.8002 | -94.123
10.8005 | -94.1235
10.8008 | -94.1233
10.8011 | -94.1231
10.8014 | -94.123
10.8017 | -94.1232
10.8021 | -94.1239
10.8024 | -94.1241
10.8027 | -94.1237
10.803 | -94.1235
10.8033 | -94.1234
10.8036 | -94.1231
10.8039 | -94.1228
10.8042 | -94.1225
10.8045 | -94.122
10.8049 | -94.1221
10.8052 | -94.122
10.8055 | -94.1219
10.8058 | -94.1216
10.8061 | -94.1216
10.8064 | -94.1217
10.8067 | -94.1218
10.807 | -94.122
10.8073 | -94.1216
10.8076 | -94.1211
10.808 | -94.121
10.8083 | -94.121
10.8086 | -94.121
10.8089 | -94.1211
10.8092 | -94.1212
10.8095 | -94.1214
10.8098 | -94.1212
10.8101 | -94.121
10.8104 | -94.1211
10.8108 | -94.1213
10.8111 | -94.1213
10.8114 | -94.1213
10.8117 | -94.1212
10.812 | -94.1208
10.8123 | -94.1207
10.8126 | -94.121
10.8129 | -94.121
10.8132 | -94.1213
10.8135 | -94.1214
10.8139 | -94.1213
10.8142 | -94.1213
10.8145 | -94.121
10.8148 | -94.1195
10.8151 | -94.1184
10.8154 | -94.1183
10.8157 | -94.1182
10.816 | -94.1181
10.8163 | -94.1177
10.8167 | -94.1176
10.817 | -94.1175
10.8173 | -94.1177
10.8176 | -94.1177
10.8179 | -94.1176
10.8182 | -94.1176
10.8185 | -94.1174
10.8188 | -94.1173
10.8191 | -94.1172
10.8194 | -94.1167
10.8198 | -94.1168
10.8201 | -94.1168
10.8204 | -94.117
10.8207 | -94.1172
10.821 | -94.1169
10.8213 | -94.1167
10.8216 | -94.1166
10.8219 | -94.1166
10.8222 | -94.1172
10.8226 | -94.1181
10.8229 | -94.118
10.8232 | -94.1174
10.8235 | -94.1178
10.8238 | -94.1185
10.8241 | -94.1187
10.8244 | -94.1186
10.8247 | -94.1182
10.825 | -94.118
10.8253 | -94.1181
10.8257 | -94.118
10.826 | -94.1179
10.8263 | -94.1175
10.8266 | -94.1169
10.8269 | -94.1162
10.8272 | -94.116
10.8275 | -94.1158
10.8278 | -94.1157
10.8281 | -94.1156
10.8285 | -94.1154
10.8288 | -94.1153
10.8291 | -94.1152
10.8294 | -94.1152
10.8297 | -94.1152
10.83 | -94.1153
10.8303 | -94.1153
10.8306 | -94.1151
10.8309 | -94.1152
10.8312 | -94.115
10.8316 | -94.1152
10.8319 | -94.1154
10.8322 | -94.1155
10.8325 | -94.1156
10.8328 | -94.1153
10.8331 | -94.1143
10.8334 | -94.1133
10.8337 | -94.1136
10.834 | -94.1143
10.8344 | -94.1145
10.8347 | -94.1145
10.835 | -94.1144
10.8353 | -94.1143
10.8356 | -94.1148
10.8359 | -94.1146
10.8362 | -94.114
10.8365 | -94.1138
10.8368 | -94.1137
10.8371 | -94.1136
10.8375 | -94.1134
10.8378 | -94.1131
10.8381 | -94.113
10.8384 | -94.1128
10.8387 | -94.1132
10.839 | -94.1143
10.8393 | -94.115
10.8396 | -94.1152
10.8399 | -94.1154
10.8403 | -94.1153
10.8406 | -94.1152
10.8409 | -94.1148
10.8412 | -94.1144
10.8415 | -94.1141
10.8418 | -94.1136
10.8421 | -94.1132
10.8424 | -94.1132
10.8427 | -94.1129
10.843 | -94.1125
10.8434 | -94.1119
10.8437 | -94.1112
10.844 | -94.1102
10.8443 | -94.1092
10.8446 | -94.1084
10.8449 | -94.1082
10.8452 | -94.1085
10.8455 | -94.1087
10.8458 | -94.1082
10.8462 | -94.1078
10.8465 | -94.1075
10.8468 | -94.107
10.8471 | -94.1066
10.8474 | -94.1064
10.8477 | -94.1064
10.848 | -94.1062
10.8483 | -94.1059
10.8486 | -94.1058
10.8489 | -94.1054
10.8493 | -94.1053
10.8496 | -94.1052
10.8499 | -94.105
10.8502 | -94.1049
10.8505 | -94.1049
10.8508 | -94.1047
10.8511 | -94.1044
10.8514 | -94.1035
10.8517 | -94.1031
10.8521 | -94.1028
10.8524 | -94.1022
10.8527 | -94.1017
10.853 | -94.1014
10.8533 | -94.1013
10.8536 | -94.101
10.8539 | -94.1008
10.8542 | -94.1003
10.8545 | -94.1003
10.8548 | -94.1002
10.8552 | -94.0996
10.8555 | -94.0989
10.8558 | -94.0985
10.8561 | -94.0983
10.8564 | -94.098
10.8567 | -94.0976
10.857 | -94.0972
10.8573 | -94.0977
10.8576 | -94.0985
10.858 | -94.099
10.8583 | -94.0986
10.8586 | -94.0993
10.8589 | -94.0998
10.8592 | -94.0985
10.8595 | -94.0973
10.8598 | -94.0967
10.8601 | -94.0967
10.8604 | -94.0975
10.8607 | -94.0977
10.8611 | -94.0968
10.8614 | -94.096
10.8617 | -94.0958
10.862 | -94.0954
10.8623 | -94.0954
10.8626 | -94.0954
10.8629 | -94.0951
10.8632 | -94.0946
10.8635 | -94.0943
10.8639 | -94.094
10.8642 | -94.0937
10.8645 | -94.0934
10.8648 | -94.093
10.8651 | -94.0921
10.8654 | -94.0915
10.8657 | -94.0914
10.866 | -94.0912
10.8663 | -94.0907
10.8666 | -94.0906
10.867 | -94.0905
10.8673 | -94.0903
10.8676 | -94.0898
10.8679 | -94.0889
10.8682 | -94.088
10.8685 | -94.0876
10.8688 | -94.087
10.8691 | -94.0862
10.8694 | -94.0857
10.8698 | -94.0859
10.8701 | -94.0861
10.8704 | -94.0854
10.8707 | -94.0851
10.871 | -94.0848
10.8713 | -94.0844
10.8716 | -94.0844
10.8719 | -94.0842
10.8722 | -94.0837
10.8725 | -94.083
10.8729 | -94.0824
10.8732 | -94.082
10.8735 | -94.0815
10.8738 | -94.0814
10.8741 | -94.0813
10.8744 | -94.0811
10.8747 | -94.0813
10.875 | -94.0807
10.8753 | -94.0799
10.8757 | -94.0796
10.876 | -94.0794
10.8763 | -94.0795
10.8766 | -94.0798
10.8769 | -94.0797
10.8772 | -94.0794
10.8775 | -94.0791
10.8778 | -94.0796
10.8781 | -94.0799
10.8784 | -94.0796
10.8788 | -94.0793
10.8791 | -94.079
10.8794 | -94.0784
10.8797 | -94.078
10.88 | -94.078
10.8803 | -94.0778
10.8806 | -94.0773
10.8809 | -94.0765
10.8812 | -94.0761
10.8816 | -94.0757
10.8819 | -94.0758
10.8822 | -94.076
10.8825 | -94.0753
10.8828 | -94.075
10.8831 | -94.0747
10.8834 | -94.0742
10.8837 | -94.0737
10.884 | -94.0734
10.8843 | -94.0732
10.8847 | -94.0726
10.885 | -94.072
10.8853 | -94.0727
10.8856 | -94.0727
10.8859 | -94.0721
10.8862 | -94.0716
10.8865 | -94.0711
10.8868 | -94.0709
10.8871 | -94.0705
10.8875 | -94.0701
10.8878 | -94.0697
10.8881 | -94.0696
10.8884 | -94.0693
10.8887 | -94.069
10.889 | -94.0687
10.8893 | -94.0684
10.8896 | -94.0681
10.8899 | -94.0678
10.8902 | -94.0677
10.8906 | -94.0673
10.8909 | -94.067
10.8912 | -94.0668
10.8915 | -94.0666
10.8918 | -94.0664
10.8921 | -94.0659
10.8924 | -94.0652
10.8927 | -94.065
10.893 | -94.0646
10.8934 | -94.0639
10.8937 | -94.0631
10.894 | -94.0628
10.8943 | -94.0624
10.8946 | -94.0618
10.8949 | -94.0611
10.8952 | -94.0607
10.8955 | -94.0602
10.8958 | -94.0596
10.8962 | -94.0589
10.8965 | -94.0584
10.8968 | -94.0585
10.8971 | -94.0587
10.8974 | -94.0585
10.8977 | -94.0582
10.898 | -94.0578
10.8983 | -94.0573
10.8986 | -94.0569
10.8989 | -94.0565
10.8993 | -94.0563
10.8996 | -94.0564
10.8999 | -94.0561
10.9002 | -94.0557
10.9005 | -94.0551
10.9008 | -94.0543
10.9011 | -94.0538
10.9014 | -94.0534
10.9017 | -94.0527
10.902 | -94.0524
10.9024 | -94.0519
10.9027 | -94.051
10.903 | -94.0505
10.9033 | -94.0503
10.9036 | -94.0495
10.9039 | -94.0486
10.9042 | -94.0485
10.9045 | -94.048
10.9048 | -94.0473
10.9052 | -94.0469
10.9055 | -94.0466
10.9058 | -94.0462
10.9061 | -94.0456
10.9064 | -94.0455
10.9067 | -94.0453
10.907 | -94.0451
10.9073 | -94.0452
10.9076 | -94.0448
10.9079 | -94.0443
10.9083 | -94.044
10.9086 | -94.0435
10.9089 | -94.0432
10.9092 | -94.0426
10.9095 | -94.0418
10.9098 | -94.0413
10.9101 | -94.0412
10.9104 | -94.0411
10.9107 | -94.0412
10.9111 | -94.0415
10.9114 | -94.0414
10.9117 | -94.0413
10.912 | -94.0408
10.9123 | -94.0406
10.9126 | -94.0402
10.9129 | -94.0398
10.9132 | -94.0392
10.9135 | -94.0389
10.9139 | -94.0385
10.9142 | -94.038
10.9145 | -94.0376
10.9148 | -94.0371
10.9151 | -94.0367
10.9154 | -94.0363
10.9157 | -94.0357
10.916 | -94.0353
10.9163 | -94.0349
10.9166 | -94.0348
10.917 | -94.0345
10.9173 | -94.0341
10.9176 | -94.0338
10.9179 | -94.0334
10.9182 | -94.0329
10.9185 | -94.0326
10.9188 | -94.0325
10.9191 | -94.0323
10.9194 | -94.0323
10.9198 | -94.0326
10.9201 | -94.0326
10.9204 | -94.0321
10.9207 | -94.0315
10.921 | -94.0315
10.9213 | -94.0313
10.9216 | -94.0306
10.9219 | -94.0299
10.9222 | -94.0295
10.9225 | -94.0291
10.9229 | -94.029
10.9232 | -94.0288
10.9235 | -94.0283
10.9238 | -94.0281
10.9241 | -94.0277
10.9244 | -94.0274
10.9247 | -94.0269
10.925 | -94.0267
10.9253 | -94.0265
10.9257 | -94.0261
10.926 | -94.026
10.9263 | -94.0256
10.9266 | -94.0252
10.9269 | -94.0248
10.9272 | -94.0241
10.9275 | -94.0234
10.9278 | -94.0231
10.9281 | -94.0226
10.9284 | -94.0223
10.9288 | -94.022
10.9291 | -94.0216
10.9294 | -94.0213
10.9297 | -94.021
10.93 | -94.0205
10.9303 | -94.02
10.9306 | -94.0195
10.9309 | -94.0193
10.9312 | -94.0193
10.9316 | -94.0189
10.9319 | -94.0186
10.9322 | -94.0182
10.9325 | -94.0177
10.9328 | -94.018
10.9331 | -94.018
10.9334 | -94.0177
10.9337 | -94.0175
10.934 | -94.0172
10.9343 | -94.0173
10.9347 | -94.0174
10.935 | -94.017
10.9353 | -94.0166
10.9356 | -94.0163
10.9359 | -94.0161
10.9362 | -94.0159
10.9365 | -94.0154
10.9368 | -94.0153
10.9371 | -94.0148
10.9375 | -94.0147
10.9378 | -94.0145
10.9381 | -94.0137
10.9384 | -94.0134
10.9387 | -94.0132
10.939 | -94.0131
10.9393 | -94.0129
10.9396 | -94.0128
10.9399 | -94.0124
10.9402 | -94.012
10.9406 | -94.0112
10.9409 | -94.0105
10.9412 | -94.01
10.9415 | -94.0097
10.9418 | -94.0091
10.9421 | -94.0087
10.9424 | -94.0086
10.9427 | -94.0083
10.943 | -94.0079
10.9434 | -94.0073
10.9437 | -94.0069
10.944 | -94.0062
10.9443 | -94.0058
10.9446 | -94.0052
10.9449 | -94.0045
10.9452 | -94.0041
10.9455 | -94.0038
10.9458 | -94.0036
10.9461 | -94.0034
10.9465 | -94.0032
10.9468 | -94.003
10.9471 | -94.0027
10.9474 | -94.0023
10.9477 | -94.0019
10.948 | -94.0017
10.9483 | -94.0014
10.9486 | -94.0008
10.9489 | -94.0004
10.9493 | -94
10.9496 | -93.9995
10.9499 | -93.9996
10.9502 | -93.9994
10.9505 | -93.9987
10.9508 | -93.9984
10.9511 | -93.9985
10.9514 | -93.9986
10.9517 | -93.9985
10.952 | -93.9983
10.9524 | -93.9978
10.9527 | -93.9972
10.953 | -93.9968
10.9533 | -93.9964
10.9536 | -93.996
10.9539 | -93.9956
10.9542 | -93.9953
10.9545 | -93.9956
10.9548 | -93.9953
10.9552 | -93.9945
10.9555 | -93.9939
10.9558 | -93.9937
10.9561 | -93.9932
10.9564 | -93.9932
10.9567 | -93.9931
10.957 | -93.9925
10.9573 | -93.9919
10.9576 | -93.9915
10.9579 | -93.9911
10.9583 | -93.9907
10.9586 | -93.9901
10.9589 | -93.9896
10.9592 | -93.9891
10.9595 | -93.9888
10.9598 | -93.9885
10.9601 | -93.9883
10.9604 | -93.9877
10.9607 | -93.9871
10.9611 | -93.9864
10.9614 | -93.9858
10.9617 | -93.9854
10.962 | -93.985
10.9623 | -93.9847
10.9626 | -93.9846
10.9629 | -93.9845
10.9632 | -93.9844
10.9635 | -93.984
10.9638 | -93.9838
10.9642 | -93.9834
10.9645 | -93.983
10.9648 | -93.9825
10.9651 | -93.982
10.9654 | -93.9815
10.9657 | -93.981
10.966 | -93.9805
10.9663 | -93.9802
10.9666 | -93.9799
10.967 | -93.9795
10.9673 | -93.9792
10.9676 | -93.979
10.9679 | -93.9789
10.9682 | -93.9787
10.9685 | -93.9783
10.9688 | -93.9779
10.9691 | -93.9775
10.9694 | -93.9771
10.9697 | -93.9767
10.9701 | -93.9768
10.9704 | -93.9765
10.9707 | -93.9759
10.971 | -93.9755
10.9713 | -93.9755
10.9716 | -93.975
10.9719 | -93.9742
10.9722 | -93.9737
10.9725 | -93.9735
10.9729 | -93.9724
10.9732 | -93.9717
10.9735 | -93.9714
10.9738 | -93.971
10.9741 | -93.9703
10.9744 | -93.9698
10.9747 | -93.9691
10.975 | -93.9682
10.9753 | -93.9677
10.9756 | -93.9675
10.976 | -93.9671
10.9763 | -93.9668
10.9766 | -93.9664
10.9769 | -93.9662
10.9772 | -93.9659
10.9775 | -93.9655
10.9778 | -93.9651
10.9781 | -93.9649
10.9784 | -93.9643
10.9788 | -93.9632
10.9791 | -93.9627
10.9794 | -93.9625
10.9797 | -93.9623
10.98 | -93.9622
10.9803 | -93.9625
10.9806 | -93.9622
10.9809 | -93.9619
10.9812 | -93.9619
10.9815 | -93.9616
10.9819 | -93.9611
10.9822 | -93.9607
10.9825 | -93.9604
10.9828 | -93.9599
10.9831 | -93.9594
10.9834 | -93.9589
10.9837 | -93.9585
10.984 | -93.9582
10.9843 | -93.9583
10.9847 | -93.9579
10.985 | -93.9579
10.9853 | -93.9575
10.9856 | -93.9573
10.9859 | -93.957
10.9862 | -93.9561
10.9865 | -93.9557
10.9868 | -93.9555
10.9871 | -93.9551
10.9874 | -93.9546
10.9878 | -93.9543
10.9881 | -93.9542
10.9884 | -93.954
10.9887 | -93.9534
10.989 | -93.9525
10.9893 | -93.952
10.9896 | -93.9517
10.9899 | -93.9518
10.9902 | -93.9514
10.9906 | -93.9509
10.9909 | -93.9503
10.9912 | -93.9498
10.9915 | -93.9497
10.9918 | -93.9493
10.9921 | -93.9481
10.9924 | -93.9476
10.9927 | -93.9476
10.993 | -93.9474
10.9933 | -93.947
10.9937 | -93.9469
10.994 | -93.9465
10.9943 | -93.946
10.9946 | -93.9455
10.9949 | -93.9447
10.9952 | -93.9441
10.9955 | -93.9436
10.9958 | -93.9433
10.9961 | -93.943
10.9965 | -93.9417
10.9968 | -93.9405
10.9971 | -93.94
10.9974 | -93.9396
10.9977 | -93.9393
10.998 | -93.9395
10.9983 | -93.9391
10.9986 | -93.9385
10.9989 | -93.9384
10.9992 | -93.9381
10.9996 | -93.9378
10.9999 | -93.9376
11.0002 | -93.9374
11.0005 | -93.937
11.0008 | -93.9367
11.0011 | -93.9369
11.0014 | -93.9372
11.0017 | -93.9374
11.002 | -93.937
11.0024 | -93.9366
11.0027 | -93.9362
11.003 | -93.9359
11.0033 | -93.9356
11.0036 | -93.9353
11.0039 | -93.9349
11.0042 | -93.9346
11.0045 | -93.9343
11.0048 | -93.934
11.0051 | -93.9339
11.0055 | -93.9335
11.0058 | -93.9331
11.0061 | -93.9328
11.0064 | -93.932
11.0067 | -93.931
11.007 | -93.9305
11.0073 | -93.9304
11.0076 | -93.9302
11.0079 | -93.9301
11.0083 | -93.9298
11.0086 | -93.9291
11.0089 | -93.9286
11.0092 | -93.9285
11.0095 | -93.9282
11.0098 | -93.9276
11.0101 | -93.9273
11.0104 | -93.9271
11.0107 | -93.9269
11.011 | -93.9268
11.0114 | -93.9259
11.0117 | -93.9253
11.012 | -93.9251
11.0123 | -93.9247
11.0126 | -93.9238
11.0129 | -93.9234
11.0132 | -93.9235
11.0135 | -93.9234
11.0138 | -93.923
11.0142 | -93.9226
11.0145 | -93.9215
11.0148 | -93.9209
11.0151 | -93.9207
11.0154 | -93.9204
11.0157 | -93.9199
11.016 | -93.9196
11.0163 | -93.9193
11.0166 | -93.9191
11.0169 | -93.9187
11.0173 | -93.9189
11.0176 | -93.9189
11.0179 | -93.9185
11.0182 | -93.9179
11.0185 | -93.9178
11.0188 | -93.9176
11.0191 | -93.9175
11.0194 | -93.9175
11.0197 | -93.9176
11.0201 | -93.9175
11.0204 | -93.9169
11.0207 | -93.9165
11.021 | -93.9162
11.0213 | -93.9158
11.0216 | -93.9156
11.0219 | -93.9157
11.0222 | -93.9151
11.0225 | -93.9144
11.0228 | -93.9135
11.0232 | -93.9132
11.0235 | -93.9131
11.0238 | -93.9127
11.0241 | -93.9123
11.0244 | -93.912
11.0247 | -93.9116
11.025 | -93.911
11.0253 | -93.9106
11.0256 | -93.9101
11.026 | -93.9098
11.0263 | -93.9094
11.0266 | -93.909
11.0269 | -93.9088
11.0272 | -93.9085
11.0275 | -93.9081
11.0278 | -93.9081
11.0281 | -93.9077
11.0284 | -93.9067
11.0288 | -93.9063
11.0291 | -93.9064
11.0294 | -93.9063
11.0297 | -93.9059
11.03 | -93.9055
11.0303 | -93.9051
11.0306 | -93.905
11.0309 | -93.9047
11.0312 | -93.9043
11.0315 | -93.9038
11.0319 | -93.9034
11.0322 | -93.9032
11.0325 | -93.9031
11.0328 | -93.9029
11.0331 | -93.9028
11.0334 | -93.9026
11.0337 | -93.9025
11.034 | -93.9022
11.0343 | -93.9021
11.0346 | -93.9023
11.035 | -93.9024
11.0353 | -93.9027
11.0356 | -93.9027
11.0359 | -93.9024
11.0362 | -93.9023
11.0365 | -93.9021
11.0368 | -93.9013
11.0371 | -93.9002
11.0374 | -93.8993
11.0378 | -93.8988
11.0381 | -93.8981
11.0384 | -93.8975
11.0387 | -93.8969
11.039 | -93.8965
11.0393 | -93.8963
11.0396 | -93.8959
11.0399 | -93.8954
11.0402 | -93.8951
11.0405 | -93.895
11.0409 | -93.8948
11.0412 | -93.8946
11.0415 | -93.8944
11.0418 | -93.8941
11.0421 | -93.8939
11.0424 | -93.8935
11.0427 | -93.8936
11.043 | -93.8941
11.0433 | -93.8939
11.0437 | -93.8935
11.044 | -93.8934
11.0443 | -93.8936
11.0446 | -93.8934
11.0449 | -93.8932
11.0452 | -93.8933
11.0455 | -93.8938
11.0458 | -93.8936
11.0461 | -93.8933
11.0465 | -93.893
11.0468 | -93.8927
11.0471 | -93.8925
11.0474 | -93.8923
11.0477 | -93.8922
11.048 | -93.8921
11.0483 | -93.8915
11.0486 | -93.8916
11.0489 | -93.8919
11.0492 | -93.8917
11.0496 | -93.8915
11.0499 | -93.8911
11.0502 | -93.8906
11.0505 | -93.8905
11.0508 | -93.8903
11.0511 | -93.8902
11.0514 | -93.8894
11.0517 | -93.8888
11.052 | -93.8886
11.0524 | -93.8883
11.0527 | -93.8879
11.053 | -93.8875
11.0533 | -93.8872
11.0536 | -93.8869
11.0539 | -93.8867
11.0542 | -93.8865
11.0545 | -93.8862
11.0548 | -93.8859
11.0551 | -93.8856
11.0555 | -93.8854
11.0558 | -93.885
11.0561 | -93.8848
11.0564 | -93.8843
11.0567 | -93.8838
11.057 | -93.8834
11.0573 | -93.8832
11.0576 | -93.8829
11.0579 | -93.8826
11.0583 | -93.8825
11.0586 | -93.8825
11.0589 | -93.8826
11.0592 | -93.8824
11.0595 | -93.8822
11.0598 | -93.882
11.0601 | -93.8817
11.0604 | -93.8815
11.0607 | -93.8814
11.061 | -93.8812
11.0614 | -93.8808
11.0617 | -93.8808
11.062 | -93.8807
11.0623 | -93.8806
11.0626 | -93.8804
11.0629 | -93.8801
11.0632 | -93.8798
11.0635 | -93.88
11.0638 | -93.8799
11.0642 | -93.8795
11.0645 | -93.8793
11.0648 | -93.8794
11.0651 | -93.8796
11.0654 | -93.8796
11.0657 | -93.8797
11.066 | -93.8794
11.0663 | -93.8787
11.0666 | -93.8785
11.0669 | -93.8786
11.0673 | -93.8785
11.0676 | -93.8782
11.0679 | -93.8783
11.0682 | -93.8784
11.0685 | -93.8788
11.0688 | -93.8787
11.0691 | -93.8783
11.0694 | -93.878
11.0697 | -93.8777
11.0701 | -93.8772
11.0704 | -93.8767
11.0707 | -93.8765
11.071 | -93.8764
11.0713 | -93.8761
11.0716 | -93.8763
11.0719 | -93.8768
11.0722 | -93.8766
11.0725 | -93.8765
11.0728 | -93.8768
11.0732 | -93.8769
11.0735 | -93.8767
11.0738 | -93.8766
11.0741 | -93.8769
11.0744 | -93.8765
11.0747 | -93.876
11.075 | -93.8758
11.0753 | -93.8756
11.0756 | -93.8754
11.076 | -93.8753
11.0763 | -93.8752
11.0766 | -93.875
11.0769 | -93.8747
11.0772 | -93.8743
11.0775 | -93.8739
11.0778 | -93.8736
11.0781 | -93.8731
11.0784 | -93.8728
11.0787 | -93.8727
11.0791 | -93.8727
11.0794 | -93.8724
11.0797 | -93.8724
11.08 | -93.8725
11.0803 | -93.8726
11.0806 | -93.873
11.0809 | -93.8733
11.0812 | -93.8732
11.0815 | -93.8727
11.0819 | -93.8717
11.0822 | -93.8713
11.0825 | -93.8712
11.0828 | -93.8711
11.0831 | -93.8715
11.0834 | -93.8715
11.0837 | -93.8717
11.084 | -93.8715
11.0843 | -93.8711
11.0846 | -93.871
11.085 | -93.871
11.0853 | -93.8706
11.0856 | -93.8705
11.0859 | -93.8703
11.0862 | -93.87
11.0865 | -93.8698
11.0868 | -93.8695
11.0871 | -93.8692
11.0874 | -93.8689
11.0878 | -93.8687
11.0881 | -93.8688
11.0884 | -93.8688
11.0887 | -93.8686
11.089 | -93.8685
11.0893 | -93.8684
11.0896 | -93.8682
11.0899 | -93.868
11.0902 | -93.8679
11.0905 | -93.8678
11.0909 | -93.8681
11.0912 | -93.8682
11.0915 | -93.868
11.0918 | -93.8677
11.0921 | -93.8674
11.0924 | -93.867
11.0927 | -93.8667
11.093 | -93.8661
11.0933 | -93.8653
11.0937 | -93.8651
11.094 | -93.8649
11.0943 | -93.8647
11.0946 | -93.8648
11.0949 | -93.865
11.0952 | -93.8653
11.0955 | -93.8652
11.0958 | -93.8651
11.0961 | -93.865
11.0964 | -93.865
11.0968 | -93.8648
11.0971 | -93.8646
11.0974 | -93.8645
11.0977 | -93.8641
11.098 | -93.8636
11.0983 | -93.8636
11.0986 | -93.864
11.0989 | -93.8641
11.0992 | -93.864
11.0996 | -93.8638
11.0999 | -93.8639
11.1002 | -93.8639
11.1005 | -93.8638
11.1008 | -93.8641
11.1011 | -93.8642
11.1014 | -93.8643
11.1017 | -93.8642
11.102 | -93.8639
11.1023 | -93.8638
11.1027 | -93.8637
11.103 | -93.8636
11.1033 | -93.8636
11.1036 | -93.8644
11.1039 | -93.8655
11.1042 | -93.8657
11.1045 | -93.8654
11.1048 | -93.865
11.1051 | -93.865
11.1055 | -93.8652
11.1058 | -93.8653
11.1061 | -93.8652
11.1064 | -93.8655
11.1067 | -93.8655
11.107 | -93.8654
11.1073 | -93.8653
11.1076 | -93.8655
11.1079 | -93.8658
11.1082 | -93.8657
11.1086 | -93.8653
11.1089 | -93.8652
11.1092 | -93.8649
11.1095 | -93.8644
11.1098 | -93.8642
11.1101 | -93.8645
11.1104 | -93.8644
11.1107 | -93.8644
11.111 | -93.8647
11.1114 | -93.8649
11.1117 | -93.8646
11.112 | -93.8644
11.1123 | -93.8645
11.1126 | -93.8644
11.1129 | -93.864
11.1132 | -93.8639
11.1135 | -93.8638
11.1138 | -93.8637
11.1141 | -93.8636
11.1145 | -93.8634
11.1148 | -93.8633
11.1151 | -93.8631
11.1154 | -93.8627
11.1157 | -93.8629
11.116 | -93.8633
11.1163 | -93.8625
11.1166 | -93.8621
11.1169 | -93.862
11.1173 | -93.862
11.1176 | -93.8621
11.1179 | -93.8622
11.1182 | -93.8622
11.1185 | -93.8622
11.1188 | -93.8626
11.1191 | -93.8624
11.1194 | -93.8625
11.1197 | -93.8627
11.12 | -93.8627
11.1204 | -93.8627
11.1207 | -93.8628
11.121 | -93.8628
11.1213 | -93.8624
11.1216 | -93.8624
11.1219 | -93.8632
11.1222 | -93.8638
11.1225 | -93.8642
11.1228 | -93.8643
11.1232 | -93.8645
11.1235 | -93.8648
11.1238 | -93.8649
11.1241 | -93.8647
11.1244 | -93.8647
11.1247 | -93.8647
11.125 | -93.8647
11.1253 | -93.8648
11.1256 | -93.8647
11.1259 | -93.8643
11.1263 | -93.8645
11.1266 | -93.8648
11.1269 | -93.8653
11.1272 | -93.8658
11.1275 | -93.8659
11.1278 | -93.8658
11.1281 | -93.8657
11.1284 | -93.8662
11.1287 | -93.8669
11.1291 | -93.8673
11.1294 | -93.8673
11.1297 | -93.8667
11.13 | -93.8665
11.1303 | -93.8667
11.1306 | -93.8672
11.1309 | -93.8671
11.1312 | -93.867
11.1315 | -93.867
11.1318 | -93.8667
11.1322 | -93.8665
11.1325 | -93.8666
11.1328 | -93.8665
11.1331 | -93.8667
11.1334 | -93.8668
11.1337 | -93.8668
11.134 | -93.8667
11.1343 | -93.8665
11.1346 | -93.8668
11.135 | -93.8673
11.1353 | -93.8679
11.1356 | -93.8681
11.1359 | -93.8681
11.1362 | -93.8686
11.1365 | -93.8692
11.1368 | -93.869
11.1371 | -93.8694
11.1374 | -93.8697
11.1377 | -93.8697
11.1381 | -93.8702
11.1384 | -93.871
11.1387 | -93.8714
11.139 | -93.8714
11.1393 | -93.8714
11.1396 | -93.8715
11.1399 | -93.8716
11.1402 | -93.8716
11.1405 | -93.8717
11.1409 | -93.8718
11.1412 | -93.8719
11.1415 | -93.8719
11.1418 | -93.872
11.1421 | -93.8725
11.1424 | -93.8728
11.1427 | -93.8728
11.143 | -93.8729
11.1433 | -93.8733
11.1436 | -93.8736
11.144 | -93.8738
11.1443 | -93.8737
11.1446 | -93.8737
11.1449 | -93.8739
11.1452 | -93.874
11.1455 | -93.8741
11.1458 | -93.8748
11.1461 | -93.875
11.1464 | -93.8753
11.1468 | -93.8756
11.1471 | -93.8758
11.1474 | -93.8761
11.1477 | -93.8763
11.148 | -93.8763
11.1483 | -93.8766
11.1486 | -93.8767
11.1489 | -93.8766
11.1492 | -93.8766
11.1495 | -93.8763
11.1499 | -93.8763
11.1502 | -93.8767
11.1505 | -93.8773
11.1508 | -93.8777
11.1511 | -93.8781
11.1514 | -93.8782
11.1517 | -93.878
11.152 | -93.8782
11.1523 | -93.8783
11.1527 | -93.8779
11.153 | -93.8777
11.1533 | -93.8779
11.1536 | -93.8782
11.1539 | -93.8783
11.1542 | -93.8782
11.1545 | -93.8783
11.1548 | -93.879
11.1551 | -93.8792
11.1554 | -93.8794
11.1558 | -93.8797
11.1561 | -93.8802
11.1564 | -93.8803
11.1567 | -93.8804
11.157 | -93.8806
11.1573 | -93.8807
11.1576 | -93.881
11.1579 | -93.8813
11.1582 | -93.8817
11.1586 | -93.882
11.1589 | -93.8823
11.1592 | -93.8823
11.1595 | -93.8824
11.1598 | -93.8824
11.1601 | -93.8825
11.1604 | -93.8827
11.1607 | -93.8829
11.161 | -93.8829
11.1613 | -93.8835
11.1617 | -93.8839
11.162 | -93.8841
11.1623 | -93.8845
11.1626 | -93.8848
11.1629 | -93.8848
11.1632 | -93.8849
11.1635 | -93.8851
11.1638 | -93.8859
11.1641 | -93.886
11.1645 | -93.8864
11.1648 | -93.8863
11.1651 | -93.8862
11.1654 | -93.8864
11.1657 | -93.8864
11.166 | -93.8864
11.1663 | -93.8864
11.1666 | -93.8864
11.1669 | -93.8867
11.1672 | -93.8874
11.1676 | -93.888
11.1679 | -93.8884
11.1682 | -93.8887
11.1685 | -93.8891
11.1688 | -93.8897
11.1691 | -93.8903
11.1694 | -93.8905
11.1697 | -93.8906
11.17 | -93.8906
11.1704 | -93.8907
11.1707 | -93.891
11.171 | -93.8914
11.1713 | -93.8916
11.1716 | -93.8918
11.1719 | -93.892
11.1722 | -93.8922
11.1725 | -93.8923
11.1728 | -93.8932
11.1731 | -93.8935
11.1735 | -93.894
11.1738 | -93.8942
11.1741 | -93.8947
11.1744 | -93.8948
11.1747 | -93.8949
11.175 | -93.8957
11.1753 | -93.8965
11.1756 | -93.8968
11.1759 | -93.897
11.1763 | -93.8973
11.1766 | -93.8976
11.1769 | -93.8984
11.1772 | -93.899
11.1775 | -93.8992
11.1778 | -93.8995
11.1781 | -93.8998
11.1784 | -93.9001
11.1787 | -93.9003
11.1791 | -93.9007
11.1794 | -93.901
11.1797 | -93.9014
11.18 | -93.9016
11.1803 | -93.9018
11.1806 | -93.9028
11.1809 | -93.903
11.1812 | -93.9029
11.1815 | -93.9037
11.1818 | -93.9037
11.1822 | -93.9037
11.1825 | -93.9042
11.1828 | -93.9047
11.1831 | -93.9053
11.1834 | -93.9056
11.1837 | -93.9055
11.184 | -93.9057
11.1843 | -93.9061
11.1846 | -93.9064
11.185 | -93.9066
11.1853 | -93.9067
11.1856 | -93.907
11.1859 | -93.9071
11.1862 | -93.9074
11.1865 | -93.908
11.1868 | -93.9086
11.1871 | -93.9091
11.1874 | -93.9091
11.1877 | -93.9093
11.1881 | -93.9095
11.1884 | -93.9092
11.1887 | -93.9094
11.189 | -93.9097
11.1893 | -93.9102
11.1896 | -93.9107
11.1899 | -93.9116
11.1902 | -93.9121
11.1905 | -93.9129
11.1908 | -93.9132
11.1912 | -93.9136
11.1915 | -93.9139
11.1918 | -93.9143
11.1921 | -93.9147
11.1924 | -93.9154
11.1927 | -93.9158
11.193 | -93.9164
11.1933 | -93.917
11.1936 | -93.9171
11.194 | -93.9174
11.1943 | -93.918
11.1946 | -93.9181
11.1949 | -93.9182
11.1952 | -93.9184
11.1955 | -93.9182
11.1958 | -93.9185
11.1961 | -93.9189
11.1964 | -93.9192
11.1968 | -93.9195
11.1971 | -93.9197
11.1974 | -93.9203
11.1977 | -93.921
11.198 | -93.9215
11.1983 | -93.922
11.1986 | -93.9222
11.1989 | -93.9223
11.1992 | -93.9224
11.1995 | -93.9229
11.1999 | -93.9228
11.2002 | -93.9228
11.2005 | -93.9235
11.2008 | -93.9243
11.2011 | -93.9248
11.2014 | -93.9252
11.2017 | -93.9257
11.202 | -93.926
11.2023 | -93.9264
11.2027 | -93.9269
11.203 | -93.9272
11.2033 | -93.9275
11.2036 | -93.9281
11.2039 | -93.9289
11.2042 | -93.9294
11.2045 | -93.9296
11.2048 | -93.9296
11.2051 | -93.9298
11.2054 | -93.9302
11.2058 | -93.9308
11.2061 | -93.931
11.2064 | -93.9309
11.2067 | -93.9312
11.207 | -93.9314
11.2073 | -93.9317
11.2076 | -93.9324
11.2079 | -93.9329
11.2082 | -93.9334
11.2086 | -93.9342
11.2089 | -93.9347
11.2092 | -93.935
11.2095 | -93.9354
11.2098 | -93.9355
11.2101 | -93.9357
11.2104 | -93.9362
11.2107 | -93.9369
11.211 | -93.9374
11.2113 | -93.9379
11.2117 | -93.9383
11.212 | -93.9385
11.2123 | -93.9388
11.2126 | -93.9391
11.2129 | -93.9394
11.2132 | -93.9399
11.2135 | -93.94
11.2138 | -93.9403
11.2141 | -93.9409
11.2145 | -93.9419
11.2148 | -93.9427
11.2151 | -93.9431
11.2154 | -93.9434
11.2157 | -93.9439
11.216 | -93.9446
11.2163 | -93.9449
11.2166 | -93.945
11.2169 | -93.945
11.2172 | -93.9453
11.2176 | -93.9463
11.2179 | -93.9467
11.2182 | -93.9471
11.2185 | -93.9476
11.2188 | -93.9486
11.2191 | -93.9497
11.2194 | -93.9502
11.2197 | -93.9506
11.22 | -93.951
11.2204 | -93.9515
11.2207 | -93.9521
11.221 | -93.9525
11.2213 | -93.9523
11.2216 | -93.9526
11.2219 | -93.953
11.2222 | -93.9534
11.2225 | -93.9538
11.2228 | -93.9544
11.2231 | -93.9549
11.2235 | -93.9556
11.2238 | -93.9562
11.2241 | -93.9567
11.2244 | -93.9572
11.2247 | -93.9578
11.225 | -93.9584
11.2253 | -93.9589
11.2256 | -93.9597
11.2259 | -93.9604
11.2263 | -93.9605
11.2266 | -93.9614
11.2269 | -93.9621
11.2272 | -93.9627
11.2275 | -93.9628
11.2278 | -93.9632
11.2281 | -93.9638
11.2284 | -93.9646
11.2287 | -93.9653
11.229 | -93.9659
11.2294 | -93.9663
11.2297 | -93.9664
11.23 | -93.9669
11.2303 | -93.9676
11.2306 | -93.9685
11.2309 | -93.9692
11.2312 | -93.9695
11.2315 | -93.97
11.2318 | -93.9705
11.2322 | -93.9712
11.2325 | -93.9715
11.2328 | -93.9716
11.2331 | -93.9722
11.2334 | -93.9726
11.2337 | -93.9732
11.234 | -93.9738
11.2343 | -93.9741
11.2346 | -93.9743
11.2349 | -93.9743
11.2353 | -93.9749
11.2356 | -93.9757
11.2359 | -93.9762
11.2362 | -93.9768
11.2365 | -93.9774
11.2368 | -93.978
11.2371 | -93.9787
11.2374 | -93.9794
11.2377 | -93.9798
11.2381 | -93.9805
11.2384 | -93.9812
11.2387 | -93.9818
11.239 | -93.9823
11.2393 | -93.9829
11.2396 | -93.9839
11.2399 | -93.9845
11.2402 | -93.9851
11.2405 | -93.9857
11.2408 | -93.9862
11.2412 | -93.9867
11.2415 | -93.9872
11.2418 | -93.988
11.2421 | -93.9887
11.2424 | -93.9892
11.2427 | -93.9896
11.243 | -93.99
11.2433 | -93.9903
11.2436 | -93.9909
11.244 | -93.9916
11.2443 | -93.9926
11.2446 | -93.9933
11.2449 | -93.9942
11.2452 | -93.9951
11.2455 | -93.9958
11.2458 | -93.9964
11.2461 | -93.9969
11.2464 | -93.9974
11.2467 | -93.9977
11.2471 | -93.9978
11.2474 | -93.9984
11.2477 | -93.9988
11.248 | -94
11.2483 | -94.0007
11.2486 | -94.0013
11.2489 | -94.002
11.2492 | -94.0025
11.2495 | -94.0031
11.2499 | -94.0036
11.2502 | -94.0042
11.2505 | -94.0048
11.2508 | -94.0053
11.2511 | -94.0057
11.2514 | -94.0064
11.2517 | -94.0077
11.252 | -94.0083
11.2523 | -94.0084
11.2526 | -94.0091
11.253 | -94.01
11.2533 | -94.0111
11.2536 | -94.012
11.2539 | -94.0128
11.2542 | -94.0132
11.2545 | -94.0132
11.2548 | -94.0136
11.2551 | -94.014
11.2554 | -94.0143
11.2558 | -94.0149
11.2561 | -94.0155
11.2564 | -94.016
11.2567 | -94.0165
11.257 | -94.0168
11.2573 | -94.0172
11.2576 | -94.0182
11.2579 | -94.0194
11.2582 | -94.0204
11.2585 | -94.0215
11.2589 | -94.0223
11.2592 | -94.0227
11.2595 | -94.0232
11.2598 | -94.0239
11.2601 | -94.0244
11.2604 | -94.025
11.2607 | -94.0255
11.261 | -94.0254
11.2613 | -94.026
11.2617 | -94.0266
11.262 | -94.0275
11.2623 | -94.0286
11.2626 | -94.0291
11.2629 | -94.0293
11.2632 | -94.0298
11.2635 | -94.0305
11.2638 | -94.031
11.2641 | -94.0314
11.2644 | -94.032
11.2648 | -94.0327
11.2651 | -94.0337
11.2654 | -94.0347
11.2657 | -94.0355
11.266 | -94.0364
11.2663 | -94.0369
11.2666 | -94.0375
11.2669 | -94.0381
11.2672 | -94.0388
11.2676 | -94.0397
11.2679 | -94.0401
11.2682 | -94.0407
11.2685 | -94.0415
11.2688 | -94.0421
11.2691 | -94.0429
11.2694 | -94.0437
11.2697 | -94.0445
11.27 | -94.0451
11.2703 | -94.0457
11.2707 | -94.0461
11.271 | -94.0465
11.2713 | -94.0472
11.2716 | -94.0477
11.2719 | -94.0483
11.2722 | -94.0489
11.2725 | -94.0495
11.2728 | -94.0503
11.2731 | -94.0508
11.2735 | -94.0513
11.2738 | -94.0521
11.2741 | -94.0528
11.2744 | -94.0534
11.2747 | -94.054
11.275 | -94.0547
11.2753 | -94.0553
11.2756 | -94.056
11.2759 | -94.0561
11.2762 | -94.0568
11.2766 | -94.0577
11.2769 | -94.0582
11.2772 | -94.0589
11.2775 | -94.0598
11.2778 | -94.0605
11.2781 | -94.061
11.2784 | -94.0619
11.2787 | -94.0627
11.279 | -94.0633
11.2794 | -94.0639
11.2797 | -94.0645
11.28 | -94.0656
11.2803 | -94.0661
11.2806 | -94.0665
11.2809 | -94.067
11.2812 | -94.0676
11.2815 | -94.0682
11.2818 | -94.0689
11.2821 | -94.0693
11.2825 | -94.0699
11.2828 | -94.071
11.2831 | -94.0723
11.2834 | -94.0729
11.2837 | -94.0736
11.284 | -94.0745
11.2843 | -94.0753
11.2846 | -94.0758
11.2849 | -94.0763
11.2853 | -94.0772
11.2856 | -94.0779
11.2859 | -94.0785
11.2862 | -94.079
11.2865 | -94.0798
11.2868 | -94.0805
11.2871 | -94.0813
11.2874 | -94.0821
11.2877 | -94.0833
11.288 | -94.0844
11.2884 | -94.085
11.2887 | -94.0856
11.289 | -94.0864
11.2893 | -94.087
11.2896 | -94.0876
11.2899 | -94.0884
11.2902 | -94.089
11.2905 | -94.0896
11.2908 | -94.0902
11.2912 | -94.091
11.2915 | -94.0917
11.2918 | -94.0923
11.2921 | -94.0928
11.2924 | -94.0934
11.2927 | -94.094
11.293 | -94.0947
11.2933 | -94.0954
11.2936 | -94.0964
11.2939 | -94.0971
11.2943 | -94.0978
11.2946 | -94.0985
11.2949 | -94.0993
11.2952 | -94.1
11.2955 | -94.1005
11.2958 | -94.1012
11.2961 | -94.1019
11.2964 | -94.1025
11.2967 | -94.1028
11.2971 | -94.1034
11.2974 | -94.1045
11.2977 | -94.1057
11.298 | -94.1066
11.2983 | -94.1067
11.2986 | -94.1071
11.2989 | -94.1079
11.2992 | -94.1087
11.2995 | -94.1095
11.2998 | -94.1103
11.3002 | -94.1115
11.3005 | -94.1125
11.3008 | -94.1135
11.3011 | -94.1149
11.3014 | -94.1161
11.3017 | -94.1167
11.302 | -94.1171
11.3023 | -94.1178
11.3026 | -94.1186
11.303 | -94.1194
11.3033 | -94.1202
11.3036 | -94.1207
11.3039 | -94.1212
11.3042 | -94.1216
11.3045 | -94.1223
11.3048 | -94.1237
11.3051 | -94.1251
11.3054 | -94.1259
11.3057 | -94.1267
11.3061 | -94.1273
11.3064 | -94.128
11.3067 | -94.1287
11.307 | -94.1297
11.3073 | -94.1305
11.3076 | -94.1312
11.3079 | -94.1321
11.3082 | -94.1328
11.3085 | -94.1337
11.3089 | -94.1343
11.3092 | -94.1352
11.3095 | -94.136
11.3098 | -94.1369
11.3101 | -94.1379
11.3104 | -94.1383
11.3107 | -94.1385
11.311 | -94.1392
11.3113 | -94.1405
11.3117 | -94.1418
11.312 | -94.1426
11.3123 | -94.1433
11.3126 | -94.1437
11.3129 | -94.1446
11.3132 | -94.1454
11.3135 | -94.1456
11.3138 | -94.1462
11.3141 | -94.1469
11.3144 | -94.1474
11.3148 | -94.1481
11.3151 | -94.1488
11.3154 | -94.1502
11.3157 | -94.1513
11.316 | -94.1521
11.3163 | -94.1529
11.3166 | -94.1537
11.3169 | -94.1547
11.3172 | -94.1555
11.3176 | -94.156
11.3179 | -94.1568
11.3182 | -94.1577
11.3185 | -94.1589
11.3188 | -94.1598
11.3191 | -94.1607
11.3194 | -94.1613
11.3197 | -94.1619
11.32 | -94.1627
11.3203 | -94.1636
11.3207 | -94.1647
11.321 | -94.1656
11.3213 | -94.1662
11.3216 | -94.1668
11.3219 | -94.1676
11.3222 | -94.1684
11.3225 | -94.1689
11.3228 | -94.1693
11.3231 | -94.1693
11.3234 | -94.1698
11.3238 | -94.1708
11.3241 | -94.1714
11.3244 | -94.1718
11.3247 | -94.1726
11.325 | -94.1736
11.3253 | -94.1743
11.3256 | -94.1751
11.3259 | -94.1759
11.3262 | -94.1767
11.3266 | -94.1772
11.3269 | -94.1778
11.3272 | -94.1786
11.3275 | -94.1793
11.3278 | -94.1798
11.3281 | -94.1804
11.3284 | -94.181
11.3287 | -94.1816
11.329 | -94.1824
11.3293 | -94.1831
11.3297 | -94.1835
11.33 | -94.184
11.3303 | -94.185
11.3306 | -94.186
11.3309 | -94.187
11.3312 | -94.1873
11.3315 | -94.1871
11.3318 | -94.1872
11.3321 | -94.1878
11.3325 | -94.1887
11.3328 | -94.1896
11.3331 | -94.1905
11.3334 | -94.1913
11.3337 | -94.1919
11.334 | -94.1929
11.3343 | -94.1936
11.3346 | -94.1941
11.3349 | -94.1949
11.3353 | -94.1957
11.3356 | -94.1962
11.3359 | -94.1971
11.3362 | -94.1981
11.3365 | -94.1989
11.3368 | -94.1996
11.3371 | -94.2002
11.3374 | -94.2009
11.3377 | -94.2017
11.338 | -94.2024
11.3384 | -94.2032
11.3387 | -94.2039
11.339 | -94.2044
11.3393 | -94.2048
11.3396 | -94.2055
11.3399 | -94.2061
11.3402 | -94.2063
11.3405 | -94.2063
11.3408 | -94.2068
11.3412 | -94.2076
11.3415 | -94.2083
11.3418 | -94.2095
11.3421 | -94.2107
11.3424 | -94.2116
11.3427 | -94.2125
11.343 | -94.2127
11.3433 | -94.2135
11.3436 | -94.2143
11.3439 | -94.2149
11.3443 | -94.2155
11.3446 | -94.2163
11.3449 | -94.217
11.3452 | -94.2178
11.3455 | -94.2186
11.3458 | -94.2193
11.3461 | -94.2201
11.3464 | -94.2208
11.3467 | -94.2218
11.3471 | -94.2226
11.3474 | -94.2236
11.3477 | -94.2246
11.348 | -94.2252
11.3483 | -94.2257
11.3486 | -94.2267
11.3489 | -94.228
11.3492 | -94.2287
11.3495 | -94.2295
11.3498 | -94.2305
11.3502 | -94.2313
11.3505 | -94.2319
11.3508 | -94.233
11.3511 | -94.2341
11.3514 | -94.2346
11.3517 | -94.2352
11.352 | -94.2355
11.3523 | -94.2361
11.3526 | -94.2368
11.353 | -94.2376
11.3533 | -94.2384
11.3536 | -94.2391
11.3539 | -94.2404
11.3542 | -94.2413
11.3545 | -94.2416
11.3548 | -94.2416
11.3551 | -94.2423
11.3554 | -94.2433
11.3557 | -94.2442
11.3561 | -94.2452
11.3564 | -94.2464
11.3567 | -94.2468
11.357 | -94.2475
11.3573 | -94.2483
11.3576 | -94.2493
11.3579 | -94.2499
11.3582 | -94.2507
11.3585 | -94.2516
11.3589 | -94.2523
11.3592 | -94.2528
11.3595 | -94.2536
11.3598 | -94.2545
11.3601 | -94.2552
11.3604 | -94.2559
11.3607 | -94.2565
11.361 | -94.2572
11.3613 | -94.2581
11.3616 | -94.2591
11.362 | -94.2599
11.3623 | -94.2606
11.3626 | -94.2614
11.3629 | -94.2623
11.3632 | -94.2628
11.3635 | -94.2634
11.3638 | -94.2643
11.3641 | -94.2655
11.3644 | -94.2664
11.3648 | -94.2672
11.3651 | -94.2678
11.3654 | -94.2683
11.3657 | -94.2693
11.366 | -94.2697
11.3663 | -94.2702
11.3666 | -94.2708
11.3669 | -94.2715
11.3672 | -94.272
11.3675 | -94.2725
11.3679 | -94.2732
11.3682 | -94.2738
11.3685 | -94.2746
11.3688 | -94.2753
11.3691 | -94.2759
11.3694 | -94.2768
11.3697 | -94.277
11.37 | -94.2782
11.3703 | -94.279
11.3707 | -94.2796
11.371 | -94.2803
11.3713 | -94.2808
11.3716 | -94.2812
11.3719 | -94.2818
11.3722 | -94.2825
11.3725 | -94.2831
11.3728 | -94.2839
11.3731 | -94.2847
11.3734 | -94.2858
11.3738 | -94.2862
11.3741 | -94.2868
11.3744 | -94.2875
11.3747 | -94.2882
11.375 | -94.2892
11.3753 | -94.2903
11.3756 | -94.2913
11.3759 | -94.2922
11.3762 | -94.2928
11.3766 | -94.2933
11.3769 | -94.2939
11.3772 | -94.2947
11.3775 | -94.2956
11.3778 | -94.2964
11.3781 | -94.2971
11.3784 | -94.2976
11.3787 | -94.2977
11.379 | -94.298
11.3793 | -94.2985
11.3797 | -94.2993
11.38 | -94.3002
11.3803 | -94.2997
11.3806 | -94.2996
11.3809 | -94.2998
11.3812 | -94.3005
11.3815 | -94.3015
11.3818 | -94.3024
11.3821 | -94.3028
11.3825 | -94.3035
11.3828 | -94.3043
11.3831 | -94.3051
11.3834 | -94.3057
11.3837 | -94.3065
11.384 | -94.3077
11.3843 | -94.3085
11.3846 | -94.3092
11.3849 | -94.3103
11.3852 | -94.3121
11.3856 | -94.3135
11.3859 | -94.3145
11.3862 | -94.3153
11.3865 | -94.3158
11.3868 | -94.3164
11.3871 | -94.3175
11.3874 | -94.3188
11.3877 | -94.3196
11.388 | -94.3206
11.3884 | -94.3213
11.3887 | -94.3221
11.389 | -94.3228
11.3893 | -94.3232
11.3896 | -94.3239
11.3899 | -94.3248
11.3902 | -94.3253
11.3905 | -94.3256
11.3908 | -94.3264
11.3911 | -94.3274
11.3915 | -94.3282
11.3918 | -94.3289
11.3921 | -94.3297
11.3924 | -94.3304
11.3927 | -94.3311
11.393 | -94.3321
11.3933 | -94.333
11.3936 | -94.3341
11.3939 | -94.335
11.3943 | -94.3352
11.3946 | -94.3357
11.3949 | -94.3363
11.3952 | -94.3371
11.3955 | -94.3377
11.3958 | -94.3384
11.3961 | -94.3391
11.3964 | -94.3393
11.3967 | -94.3402
11.397 | -94.3409
11.3974 | -94.3416
11.3977 | -94.3422
11.398 | -94.3429
11.3983 | -94.3434
11.3986 | -94.344
11.3989 | -94.3447
11.3992 | -94.3452
11.3995 | -94.3457
11.3998 | -94.3461
11.4002 | -94.3466
11.4005 | -94.3469
11.4008 | -94.3474
11.4011 | -94.3479
11.4014 | -94.3483
11.4017 | -94.349
11.402 | -94.3497
11.4023 | -94.3504
11.4026 | -94.3511
11.4029 | -94.352
11.4033 | -94.3525
11.4036 | -94.3531
11.4039 | -94.3539
11.4042 | -94.3546
11.4045 | -94.3554
11.4048 | -94.3563
11.4051 | -94.3571
11.4054 | -94.3579
11.4057 | -94.3579
11.4061 | -94.3583
11.4064 | -94.3588
11.4067 | -94.3596
11.407 | -94.3605
11.4073 | -94.3612
11.4076 | -94.3619
11.4079 | -94.3625
11.4082 | -94.3631
11.4085 | -94.3639
11.4088 | -94.3645
11.4092 | -94.3656
11.4095 | -94.3664
11.4098 | -94.3668
11.4101 | -94.3676
11.4104 | -94.3682
11.4107 | -94.3683
11.411 | -94.369
11.4113 | -94.3695
11.4116 | -94.3697
11.412 | -94.3701
11.4123 | -94.3707
11.4126 | -94.3718
11.4129 | -94.3722
11.4132 | -94.3723
11.4135 | -94.3727
11.4138 | -94.3734
11.4141 | -94.3742
11.4144 | -94.3746
11.4147 | -94.3747
11.4151 | -94.3752
11.4154 | -94.3756
11.4157 | -94.3764
11.416 | -94.3771
11.4163 | -94.3778
11.4166 | -94.3788
11.4169 | -94.3793
11.4172 | -94.3799
11.4175 | -94.3809
11.4179 | -94.3818
11.4182 | -94.3824
11.4185 | -94.3833
11.4188 | -94.3841
11.4191 | -94.3847
11.4194 | -94.3855
11.4197 | -94.3861
11.42 | -94.3872
11.4203 | -94.3886
11.4206 | -94.3893
11.421 | -94.3895
11.4213 | -94.3902
11.4216 | -94.391
11.4219 | -94.3916
11.4222 | -94.3928
11.4225 | -94.3946
11.4228 | -94.3961
11.4231 | -94.3967
11.4234 | -94.397
11.4238 | -94.3985
11.4241 | -94.3997
11.4244 | -94.4011
11.4247 | -94.4027
11.425 | -94.4036
11.4253 | -94.4043
11.4256 | -94.4048
11.4259 | -94.4054
11.4262 | -94.4059
11.4265 | -94.4064
11.4269 | -94.407
11.4272 | -94.4074
11.4275 | -94.4077
11.4278 | -94.4081
11.4281 | -94.4088
11.4284 | -94.4095
11.4287 | -94.4101
11.429 | -94.4107
11.4293 | -94.4114
11.4297 | -94.4119
11.43 | -94.4123
11.4303 | -94.4127
11.4306 | -94.4133
11.4309 | -94.4139
11.4312 | -94.4145
11.4315 | -94.4157
11.4318 | -94.4168
11.4321 | -94.4174
11.4324 | -94.4175
11.4328 | -94.4184
11.4331 | -94.419
11.4334 | -94.4197
11.4337 | -94.4204
11.434 | -94.4211
11.4343 | -94.4217
11.4346 | -94.4225
11.4349 | -94.4232
11.4352 | -94.4238
11.4356 | -94.4243
11.4359 | -94.4245
11.4362 | -94.4246
11.4365 | -94.425
11.4368 | -94.4256
11.4371 | -94.4261
11.4374 | -94.4267
11.4377 | -94.4277
11.438 | -94.4288
11.4383 | -94.4293
11.4387 | -94.4298
11.439 | -94.4304
11.4393 | -94.4306
11.4396 | -94.4315
11.4399 | -94.4324
11.4402 | -94.434
11.4405 | -94.4359
11.4408 | -94.4367
11.4411 | -94.4372
11.4415 | -94.4376
11.4418 | -94.4385
11.4421 | -94.4394
11.4424 | -94.4402
11.4427 | -94.4408
11.443 | -94.4414
11.4433 | -94.442
11.4436 | -94.4425
11.4439 | -94.4432
11.4442 | -94.4438
11.4446 | -94.4443
11.4449 | -94.4454
11.4452 | -94.4465
11.4455 | -94.4472
11.4458 | -94.448
11.4461 | -94.4486
11.4464 | -94.449
11.4467 | -94.4497
11.447 | -94.4501
11.4474 | -94.4505
11.4477 | -94.4509
11.448 | -94.4515
11.4483 | -94.4524
11.4486 | -94.4531
11.4489 | -94.4536
11.4492 | -94.4543
11.4495 | -94.4547
11.4498 | -94.4553
11.4501 | -94.4562
11.4505 | -94.4568
11.4508 | -94.4574
11.4511 | -94.4582
11.4514 | -94.4591
11.4517 | -94.4596
11.452 | -94.4602
11.4523 | -94.4606
11.4526 | -94.4609
11.4529 | -94.4613
11.4533 | -94.4611
11.4536 | -94.4606
11.4539 | -94.4609
11.4542 | -94.4617
11.4545 | -94.4622
11.4548 | -94.4625
11.4551 | -94.4633
11.4554 | -94.4637
11.4557 | -94.4642
11.456 | -94.4651
11.4564 | -94.4657
11.4567 | -94.4666
11.457 | -94.4674
11.4573 | -94.4678
11.4576 | -94.468
11.4579 | -94.4685
11.4582 | -94.4692
11.4585 | -94.4697
11.4588 | -94.4699
11.4592 | -94.4703
11.4595 | -94.4703
11.4598 | -94.4708
11.4601 | -94.4713
11.4604 | -94.4718
11.4607 | -94.4724
11.461 | -94.4728
11.4613 | -94.4735
11.4616 | -94.4742
11.4619 | -94.4754
11.4623 | -94.4772
11.4626 | -94.4783
11.4629 | -94.4784
11.4632 | -94.4786
11.4635 | -94.4788
11.4638 | -94.4792
11.4641 | -94.4797
11.4644 | -94.4801
11.4647 | -94.4803
11.4651 | -94.4807
11.4654 | -94.4811
11.4657 | -94.4817
11.466 | -94.4822
11.4663 | -94.4824
11.4666 | -94.483
11.4669 | -94.4835
11.4672 | -94.4834
11.4675 | -94.4837
11.4679 | -94.4844
11.4682 | -94.4847
11.4685 | -94.4848
11.4688 | -94.4848
11.4691 | -94.4861
11.4694 | -94.4865
11.4697 | -94.4871
11.47 | -94.4878
11.4703 | -94.4884
11.4706 | -94.4894
11.471 | -94.4902
11.4713 | -94.4906
11.4716 | -94.4912
11.4719 | -94.4917
11.4722 | -94.4922
11.4725 | -94.4929
11.4728 | -94.4932
11.4731 | -94.4937
11.4734 | -94.494
11.4738 | -94.494
11.4741 | -94.4943
11.4744 | -94.4944
11.4747 | -94.4946
11.475 | -94.495
11.4753 | -94.4958
11.4756 | -94.4969
11.4759 | -94.4978
11.4762 | -94.499
11.4765 | -94.5002
11.4769 | -94.5007
11.4772 | -94.501
11.4775 | -94.5012
11.4778 | -94.5012
11.4781 | -94.5015
11.4784 | -94.5016
11.4787 | -94.502
11.479 | -94.5022
11.4793 | -94.5023
11.4796 | -94.5029
11.48 | -94.5039
11.4803 | -94.5044
11.4806 | -94.5047
11.4809 | -94.505
11.4812 | -94.5061
11.4815 | -94.5069
11.4818 | -94.5074
11.4821 | -94.5077
11.4824 | -94.5081
11.4828 | -94.5086
11.4831 | -94.5091
11.4834 | -94.5093
11.4837 | -94.5092
11.484 | -94.5094
11.4843 | -94.5095
11.4846 | -94.51
11.4849 | -94.5105
11.4852 | -94.511
11.4856 | -94.5114
11.4859 | -94.5118
11.4862 | -94.5122
11.4865 | -94.5123
11.4868 | -94.5123
11.4871 | -94.512
11.4874 | -94.5122
11.4877 | -94.5126
11.488 | -94.513
11.4883 | -94.5133
11.4887 | -94.514
11.489 | -94.5144
11.4893 | -94.5149
11.4896 | -94.5154
11.4899 | -94.5152
11.4902 | -94.5154
11.4905 | -94.516
11.4908 | -94.5163
11.4911 | -94.5168
11.4915 | -94.5175
11.4918 | -94.518
11.4921 | -94.5182
11.4924 | -94.5186
11.4927 | -94.5194
11.493 | -94.5201
11.4933 | -94.5206
11.4936 | -94.5209
11.4939 | -94.5213
11.4942 | -94.5218
11.4946 | -94.5222
11.4949 | -94.5226
11.4952 | -94.5231
11.4955 | -94.5236
11.4958 | -94.5233
11.4961 | -94.5223
11.4964 | -94.5222
11.4967 | -94.5229
11.497 | -94.5237
11.4974 | -94.5247
11.4977 | -94.5257
11.498 | -94.5264
11.4983 | -94.5268
11.4986 | -94.5263
11.4989 | -94.5266
11.4992 | -94.5273
11.4995 | -94.528
11.4998 | -94.5279
11.5001 | -94.5278
11.5005 | -94.5281
11.5008 | -94.5282
11.5011 | -94.528
11.5014 | -94.5282
11.5017 | -94.5289
11.502 | -94.5295
11.5023 | -94.53
11.5026 | -94.5307
11.5029 | -94.5313
11.5033 | -94.5316
11.5036 | -94.532
11.5039 | -94.5325
11.5042 | -94.5328
11.5045 | -94.5331
11.5048 | -94.5335
11.5051 | -94.534
11.5054 | -94.5353
11.5057 | -94.5361
11.506 | -94.5365
11.5064 | -94.5368
11.5067 | -94.5372
11.507 | -94.5374
11.5073 | -94.5378
11.5076 | -94.538
11.5079 | -94.5381
11.5082 | -94.5382
11.5085 | -94.5386
11.5088 | -94.5387
11.5092 | -94.5391
11.5095 | -94.5394
11.5098 | -94.5398
11.5101 | -94.5398
11.5104 | -94.54
11.5107 | -94.5407
11.511 | -94.5412
11.5113 | -94.5417
11.5116 | -94.5425
11.5119 | -94.5433
11.5123 | -94.5434
11.5126 | -94.5437
11.5129 | -94.5439
11.5132 | -94.5442
11.5135 | -94.5447
11.5138 | -94.5455
11.5141 | -94.5463
11.5144 | -94.5468
11.5147 | -94.5472
11.5151 | -94.5475
11.5154 | -94.5477
11.5157 | -94.5481
11.516 | -94.5485
11.5163 | -94.5487
11.5166 | -94.5486
11.5169 | -94.5488
11.5172 | -94.5486
11.5175 | -94.5489
11.5178 | -94.5492
11.5182 | -94.5499
11.5185 | -94.5504
11.5188 | -94.5507
11.5191 | -94.5509
11.5194 | -94.5512
11.5197 | -94.552
11.52 | -94.5524
11.5203 | -94.5525
11.5206 | -94.5531
11.521 | -94.5532
11.5213 | -94.5535
11.5216 | -94.5535
11.5219 | -94.5535
11.5222 | -94.5536
11.5225 | -94.5538
11.5228 | -94.5539
11.5231 | -94.5542
11.5234 | -94.5542
11.5237 | -94.5542
11.5241 | -94.5545
11.5244 | -94.5548
11.5247 | -94.5549
11.525 | -94.5551
11.5253 | -94.5557
11.5256 | -94.5559
11.5259 | -94.5561
11.5262 | -94.5564
11.5265 | -94.5569
11.5269 | -94.5573
11.5272 | -94.5574
11.5275 | -94.5576
11.5278 | -94.5578
11.5281 | -94.558
11.5284 | -94.5582
11.5287 | -94.5579
11.529 | -94.558
11.5293 | -94.5582
11.5296 | -94.5578
11.53 | -94.5572
11.5303 | -94.5574
11.5306 | -94.5577
11.5309 | -94.5582
11.5312 | -94.5583
11.5315 | -94.5583
11.5318 | -94.5581
11.5321 | -94.5583
11.5324 | -94.5584
11.5328 | -94.5583
11.5331 | -94.5584
11.5334 | -94.559
11.5337 | -94.5594
11.534 | -94.5597
11.5343 | -94.5607
11.5346 | -94.5616
11.5349 | -94.5613
11.5352 | -94.561
11.5355 | -94.561
11.5359 | -94.5616
11.5362 | -94.5624
11.5365 | -94.5628
11.5368 | -94.563
11.5371 | -94.5634
11.5374 | -94.5636
11.5377 | -94.5632
11.538 | -94.5628
11.5383 | -94.5627
11.5387 | -94.5628
11.539 | -94.5633
11.5393 | -94.5636
11.5396 | -94.5643
11.5399 | -94.5644
11.5402 | -94.5651
11.5405 | -94.5656
11.5408 | -94.5661
11.5411 | -94.5663
11.5414 | -94.5667
11.5418 | -94.5672
11.5421 | -94.5678
11.5424 | -94.5683
11.5427 | -94.5687
11.543 | -94.5689
11.5433 | -94.5691
11.5436 | -94.5694
11.5439 | -94.5698
11.5442 | -94.5698
11.5446 | -94.57
11.5449 | -94.5701
11.5452 | -94.5702
11.5455 | -94.5699
11.5458 | -94.5702
11.5461 | -94.5706
11.5464 | -94.571
11.5467 | -94.5715
11.547 | -94.5718
11.5473 | -94.572
11.5477 | -94.5722
11.548 | -94.5725
11.5483 | -94.5729
11.5486 | -94.5734
11.5489 | -94.5735
11.5492 | -94.5743
11.5495 | -94.5751
11.5498 | -94.5758
11.5501 | -94.5758
11.5505 | -94.5751
11.5508 | -94.5747
11.5511 | -94.5746
11.5514 | -94.5749
11.5517 | -94.576
11.552 | -94.5771
11.5523 | -94.5775
11.5526 | -94.5776
11.5529 | -94.5778
11.5532 | -94.5781
11.5536 | -94.5776
11.5539 | -94.578
11.5542 | -94.5782
11.5545 | -94.5776
11.5548 | -94.5774
11.5551 | -94.5776
11.5554 | -94.5778
11.5557 | -94.5784
11.556 | -94.5794
11.5564 | -94.5797
11.5567 | -94.5798
11.557 | -94.5801
11.5573 | -94.5801
11.5576 | -94.5799
11.5579 | -94.58
11.5582 | -94.5804
11.5585 | -94.5807
11.5588 | -94.5809
11.5591 | -94.5811
11.5595 | -94.5811
11.5598 | -94.5809
11.5601 | -94.581
11.5604 | -94.581
11.5607 | -94.5811
11.561 | -94.5812
11.5613 | -94.5815
11.5616 | -94.5815
11.5619 | -94.5815
11.5623 | -94.5817
11.5626 | -94.5812
11.5629 | -94.5802
11.5632 | -94.5794
11.5635 | -94.5794
11.5638 | -94.5796
11.5641 | -94.5799
11.5644 | -94.5802
11.5647 | -94.5802
11.565 | -94.5801
11.5654 | -94.58
11.5657 | -94.58
11.566 | -94.5799
11.5663 | -94.5795
11.5666 | -94.5796
11.5669 | -94.5797
11.5672 | -94.58
11.5675 | -94.5802
11.5678 | -94.5803
11.5682 | -94.5809
11.5685 | -94.5816
11.5688 | -94.5821
11.5691 | -94.5823
11.5694 | -94.5824
11.5697 | -94.5822
11.57 | -94.5817
11.5703 | -94.5821
11.5706 | -94.5824
11.5709 | -94.583
11.5713 | -94.5832
11.5716 | -94.5826
11.5719 | -94.5825
11.5722 | -94.583
11.5725 | -94.5831
11.5728 | -94.5832
11.5731 | -94.5834
11.5734 | -94.5835
11.5737 | -94.5835
11.5741 | -94.5839
11.5744 | -94.5844
11.5747 | -94.5841
11.575 | -94.5838
11.5753 | -94.5843
11.5756 | -94.5846
11.5759 | -94.5846
11.5762 | -94.5843
11.5765 | -94.5843
11.5768 | -94.5847
11.5772 | -94.585
11.5775 | -94.5851
11.5778 | -94.5851
11.5781 | -94.5853
11.5784 | -94.585
11.5787 | -94.5852
11.579 | -94.5862
11.5793 | -94.5865
11.5796 | -94.5866
11.58 | -94.5862
11.5803 | -94.5863
11.5806 | -94.5869
11.5809 | -94.5874
11.5812 | -94.5876
11.5815 | -94.5876
11.5818 | -94.5877
11.5821 | -94.5881
11.5824 | -94.5887
11.5827 | -94.5889
11.5831 | -94.5886
11.5834 | -94.5884
11.5837 | -94.5882
11.584 | -94.5884
11.5843 | -94.5889
11.5846 | -94.5887
11.5849 | -94.5886
11.5852 | -94.5886
11.5855 | -94.5886
11.5859 | -94.5886
11.5862 | -94.5888
11.5865 | -94.5888
11.5868 | -94.5889
11.5871 | -94.5891
11.5874 | -94.589
11.5877 | -94.5891
11.588 | -94.5889
11.5883 | -94.5889
11.5886 | -94.5889
11.589 | -94.5886
11.5893 | -94.5887
11.5896 | -94.5887
11.5899 | -94.5888
11.5902 | -94.5888
11.5905 | -94.5888
11.5908 | -94.5889
11.5911 | -94.5889
11.5914 | -94.5891
11.5918 | -94.5893
11.5921 | -94.5906
11.5924 | -94.5917
11.5927 | -94.5917
11.593 | -94.5914
11.5933 | -94.5913
11.5936 | -94.5917
11.5939 | -94.5914
11.5942 | -94.5914
11.5945 | -94.5916
11.5949 | -94.5928
11.5952 | -94.5936
11.5955 | -94.5936
11.5958 | -94.5934
11.5961 | -94.5933
11.5964 | -94.5933
11.5967 | -94.5933
11.597 | -94.5938
11.5973 | -94.5938
11.5977 | -94.5938
11.598 | -94.5939
11.5983 | -94.5939
11.5986 | -94.594
11.5989 | -94.594
11.5992 | -94.594
11.5995 | -94.594
11.5998 | -94.5941
11.6001 | -94.5946
11.6005 | -94.5947
11.6008 | -94.595
11.6011 | -94.5954
11.6014 | -94.5955
11.6017 | -94.5955
11.602 | -94.5955
11.6023 | -94.5953
11.6026 | -94.5954
11.6029 | -94.5957
11.6032 | -94.5956
11.6036 | -94.5957
11.6039 | -94.5961
11.6042 | -94.596
11.6045 | -94.5964
11.6048 | -94.5971
11.6051 | -94.5975
11.6054 | -94.597
11.6057 | -94.5966
11.606 | -94.5966
11.6064 | -94.5969
11.6067 | -94.597
11.607 | -94.597
11.6073 | -94.5967
11.6076 | -94.5967
11.6079 | -94.5967
11.6082 | -94.597
11.6085 | -94.5971
11.6088 | -94.597
11.6091 | -94.5973
11.6095 | -94.5977
11.6098 | -94.5975
11.6101 | -94.5974
11.6104 | -94.5974
11.6107 | -94.5973
11.611 | -94.5973
11.6113 | -94.5975
11.6116 | -94.5974
11.6119 | -94.5984
11.6122 | -94.5988
11.6126 | -94.5984
11.6129 | -94.5984
11.6132 | -94.5985
11.6135 | -94.5984
11.6138 | -94.5984
11.6141 | -94.5988
11.6144 | -94.5989
11.6147 | -94.5992
11.615 | -94.5998
11.6154 | -94.6001
11.6157 | -94.6
11.616 | -94.5999
11.6163 | -94.6
11.6166 | -94.6
11.6169 | -94.6
11.6172 | -94.5998
11.6175 | -94.5992
11.6178 | -94.599
11.6181 | -94.5993
11.6185 | -94.5994
11.6188 | -94.5994
11.6191 | -94.5984
11.6194 | -94.5975
11.6197 | -94.5978
11.62 | -94.598
11.6203 | -94.5981
11.6206 | -94.598
11.6209 | -94.5981
11.6213 | -94.5985
11.6216 | -94.5985
11.6219 | -94.5982
11.6222 | -94.5982
11.6225 | -94.5982
11.6228 | -94.5982
11.6231 | -94.5995
11.6234 | -94.6008
11.6237 | -94.6015
11.6241 | -94.6019
11.6244 | -94.6021
11.6247 | -94.6019
11.625 | -94.6017
11.6253 | -94.602
11.6256 | -94.602
11.6259 | -94.6014
11.6262 | -94.6012
11.6265 | -94.6015
11.6268 | -94.6014
11.6272 | -94.6011
11.6275 | -94.6011
11.6278 | -94.601
11.6281 | -94.601
11.6284 | -94.6014
11.6287 | -94.6013
11.629 | -94.6011
11.6293 | -94.6008
11.6296 | -94.5998
11.63 | -94.5995
11.6303 | -94.5995
11.6306 | -94.5994
11.6309 | -94.5996
11.6312 | -94.5993
11.6315 | -94.5989
11.6318 | -94.599
11.6321 | -94.5992
11.6324 | -94.5996
11.6327 | -94.5991
11.6331 | -94.5983
11.6334 | -94.5979
11.6337 | -94.5978
11.634 | -94.5977
11.6343 | -94.5979
11.6346 | -94.5978
11.6349 | -94.5976
11.6352 | -94.5975
11.6355 | -94.5978
11.6359 | -94.5978
11.6362 | -94.5975
11.6365 | -94.5975
11.6368 | -94.5974
11.6371 | -94.5974
11.6374 | -94.5972
11.6377 | -94.5968
11.638 | -94.5964
11.6383 | -94.5964
11.6386 | -94.5955
11.639 | -94.5956
11.6393 | -94.5955
11.6396 | -94.5957
11.6399 | -94.5961
11.6402 | -94.597
11.6405 | -94.5975
11.6408 | -94.5971
11.6411 | -94.5966
11.6414 | -94.5962
11.6418 | -94.5966
11.6421 | -94.5972
11.6424 | -94.5974
11.6427 | -94.5973
11.643 | -94.5971
11.6433 | -94.5968
11.6436 | -94.5966
11.6439 | -94.5965
11.6442 | -94.5962
11.6445 | -94.5961
11.6449 | -94.5962
11.6452 | -94.5958
11.6455 | -94.5959
11.6458 | -94.5966
11.6461 | -94.5971
11.6464 | -94.5969
11.6467 | -94.5964
11.647 | -94.596
11.6473 | -94.5958
11.6477 | -94.5958
11.648 | -94.5957
11.6483 | -94.5955
11.6486 | -94.5954
11.6489 | -94.596
11.6492 | -94.5964
11.6495 | -94.5964
11.6498 | -94.596
11.6501 | -94.5958
11.6504 | -94.5957
11.6508 | -94.5957
11.6511 | -94.5953
11.6514 | -94.5949
11.6517 | -94.595
11.652 | -94.595
11.6523 | -94.5948
11.6526 | -94.5948
11.6529 | -94.5946
11.6532 | -94.5945
11.6536 | -94.5945
11.6539 | -94.5945
11.6542 | -94.5947
11.6545 | -94.5946
11.6548 | -94.5943
11.6551 | -94.5943
11.6554 | -94.5942
11.6557 | -94.594
11.656 | -94.5937
11.6563 | -94.5936
11.6567 | -94.5932
11.657 | -94.5931
11.6573 | -94.5926
11.6576 | -94.5925
11.6579 | -94.5923
11.6582 | -94.5922
11.6585 | -94.5917
11.6588 | -94.5915
11.6591 | -94.5913
11.6595 | -94.5911
11.6598 | -94.5909
11.6601 | -94.5906
11.6604 | -94.5899
11.6607 | -94.589
11.661 | -94.5881
11.6613 | -94.5879
11.6616 | -94.5877
11.6619 | -94.5875
11.6622 | -94.587
11.6626 | -94.587
11.6629 | -94.5873
11.6632 | -94.5867
11.6635 | -94.5863
11.6638 | -94.5861
11.6641 | -94.5858
11.6644 | -94.5853
11.6647 | -94.5851
11.665 | -94.5848
11.6654 | -94.5846
11.6657 | -94.5846
11.666 | -94.5847
11.6663 | -94.5843
11.6666 | -94.584
11.6669 | -94.5838
11.6672 | -94.5838
11.6675 | -94.5836
11.6678 | -94.5831
11.6681 | -94.5826
11.6685 | -94.5826
11.6688 | -94.5826
11.6691 | -94.5827
11.6694 | -94.5826
11.6697 | -94.5823
11.67 | -94.5821
11.6703 | -94.5819
11.6706 | -94.5818
11.6709 | -94.5811
11.6713 | -94.5804
11.6716 | -94.58
11.6719 | -94.5797
11.6722 | -94.5794
11.6725 | -94.5791
11.6728 | -94.5786
11.6731 | -94.578
11.6734 | -94.5778
11.6737 | -94.5775
11.674 | -94.5772
11.6744 | -94.5769
11.6747 | -94.5767
11.675 | -94.5764
11.6753 | -94.5769
11.6756 | -94.5769
11.6759 | -94.5763
11.6762 | -94.5764
11.6765 | -94.5768
11.6768 | -94.5764
11.6772 | -94.576
11.6775 | -94.5757
11.6778 | -94.5758
11.6781 | -94.5757
11.6784 | -94.5753
11.6787 | -94.5748
11.679 | -94.5748
11.6793 | -94.5746
11.6796 | -94.5739
11.6799 | -94.5737
11.6803 | -94.5734
11.6806 | -94.5729
11.6809 | -94.5721
11.6812 | -94.5716
11.6815 | -94.5714
11.6818 | -94.5715
11.6821 | -94.5709
11.6824 | -94.5705
11.6827 | -94.5703
11.6831 | -94.57
11.6834 | -94.5697
11.6837 | -94.5695
11.684 | -94.5694
11.6843 | -94.5696
11.6846 | -94.5698
11.6849 | -94.5699
11.6852 | -94.57
11.6855 | -94.5699
11.6858 | -94.5697
11.6862 | -94.57
11.6865 | -94.5698
11.6868 | -94.5691
11.6871 | -94.5692
11.6874 | -94.5694
11.6877 | -94.5693
11.688 | -94.5692
11.6883 | -94.569
11.6886 | -94.569
11.689 | -94.5679
11.6893 | -94.5675
11.6896 | -94.5669
11.6899 | -94.5669
11.6902 | -94.5667
11.6905 | -94.5667
11.6908 | -94.5669
11.6911 | -94.5666
11.6914 | -94.5662
11.6917 | -94.566
11.6921 | -94.5658
11.6924 | -94.5652
11.6927 | -94.5653
11.693 | -94.565
11.6933 | -94.5649
11.6936 | -94.5645
11.6939 | -94.5644
11.6942 | -94.5637
11.6945 | -94.5633
11.6949 | -94.563
11.6952 | -94.5625
11.6955 | -94.5621
11.6958 | -94.5617
11.6961 | -94.5618
11.6964 | -94.5615
11.6967 | -94.5612
11.697 | -94.5611
11.6973 | -94.5609
11.6976 | -94.5609
11.698 | -94.5608
11.6983 | -94.5605
11.6986 | -94.5603
11.6989 | -94.5598
11.6992 | -94.5595
11.6995 | -94.5588
11.6998 | -94.558
11.7001 | -94.5577
11.7004 | -94.5571
11.7008 | -94.5565
11.7011 | -94.5564
11.7014 | -94.5563
11.7017 | -94.5563
11.702 | -94.5561
11.7023 | -94.556
11.7026 | -94.5556
11.7029 | -94.5553
11.7032 | -94.555
11.7035 | -94.5545
11.7039 | -94.5541
11.7042 | -94.5534
11.7045 | -94.5522
11.7048 | -94.5513
11.7051 | -94.551
11.7054 | -94.5506
11.7057 | -94.5498
11.706 | -94.5489
11.7063 | -94.5483
11.7067 | -94.5474
11.707 | -94.547
11.7073 | -94.5469
11.7076 | -94.5464
11.7079 | -94.545
11.7082 | -94.5442
11.7085 | -94.5439
11.7088 | -94.5438
11.7091 | -94.5436
11.7094 | -94.5435
11.7098 | -94.5438
11.7101 | -94.5437
11.7104 | -94.5435
11.7107 | -94.5429
11.711 | -94.5421
11.7113 | -94.5416
11.7116 | -94.5411
11.7119 | -94.5407
11.7122 | -94.5402
11.7126 | -94.5401
11.7129 | -94.5382
11.7132 | -94.5358
11.7135 | -94.535
11.7138 | -94.535
11.7141 | -94.5349
11.7144 | -94.5344
11.7147 | -94.5334
11.715 | -94.5327
11.7153 | -94.5325
11.7157 | -94.5323
11.716 | -94.5318
11.7163 | -94.5316
11.7166 | -94.5304
11.7169 | -94.5291
11.7172 | -94.5289
11.7175 | -94.5287
11.7178 | -94.5284
11.7181 | -94.5282
11.7185 | -94.528
11.7188 | -94.5276
11.7191 | -94.5268
11.7194 | -94.5261
11.7197 | -94.5256
11.72 | -94.5253
11.7203 | -94.5249
11.7206 | -94.5243
11.7209 | -94.5239
11.7212 | -94.5233
11.7216 | -94.522
11.7219 | -94.5217
11.7222 | -94.5209
11.7225 | -94.5201
11.7228 | -94.5197
11.7231 | -94.5193
11.7234 | -94.519
11.7237 | -94.5189
11.724 | -94.5189
11.7244 | -94.5187
11.7247 | -94.5184
11.725 | -94.5181
11.7253 | -94.5178
11.7256 | -94.5174
11.7259 | -94.5171
11.7262 | -94.5169
11.7265 | -94.5167
11.7268 | -94.5167
11.7271 | -94.5167
11.7275 | -94.5163
11.7278 | -94.5161
11.7281 | -94.516
11.7284 | -94.5158
11.7287 | -94.5152
11.729 | -94.5142
11.7293 | -94.5131
11.7296 | -94.5122
11.7299 | -94.5117
11.7303 | -94.511
11.7306 | -94.5102
11.7309 | -94.5093
11.7312 | -94.5081
11.7315 | -94.5078
11.7318 | -94.5074
11.7321 | -94.507
11.7324 | -94.5066
11.7327 | -94.5062
11.733 | -94.5055
11.7334 | -94.5049
11.7337 | -94.5044
11.734 | -94.5039
11.7343 | -94.5028
11.7346 | -94.5021
11.7349 | -94.5019
11.7352 | -94.5018
11.7355 | -94.5013
11.7358 | -94.5007
11.7362 | -94.5
11.7365 | -94.4997
11.7368 | -94.4992
11.7371 | -94.4989
11.7374 | -94.4985
11.7377 | -94.4985
11.738 | -94.4981
11.7383 | -94.4976
11.7386 | -94.4972
11.739 | -94.4969
11.7393 | -94.4965
11.7396 | -94.4964
11.7399 | -94.4963
11.7402 | -94.4955
11.7405 | -94.4947
11.7408 | -94.4941
11.7411 | -94.4935
11.7414 | -94.493
11.7417 | -94.4926
11.7421 | -94.4926
11.7424 | -94.4925
11.7427 | -94.4924
11.743 | -94.492
11.7433 | -94.4915
11.7436 | -94.4907
11.7439 | -94.4903
11.7442 | -94.4903
11.7445 | -94.4902
11.7448 | -94.4898
11.7452 | -94.4895
11.7455 | -94.4891
11.7458 | -94.4887
11.7461 | -94.4878
11.7464 | -94.4869
11.7467 | -94.4865
11.747 | -94.4861
11.7473 | -94.4856
11.7476 | -94.4852
11.748 | -94.4848
11.7483 | -94.4845
11.7486 | -94.4848
11.7489 | -94.4854
11.7492 | -94.4851
11.7495 | -94.4842
11.7498 | -94.4836
11.7501 | -94.4832
11.7504 | -94.4826
11.7508 | -94.4821
11.7511 | -94.4819
11.7514 | -94.4814
11.7517 | -94.4807
11.752 | -94.48
11.7523 | -94.4791
11.7526 | -94.4788
11.7529 | -94.4784
11.7532 | -94.4782
11.7535 | -94.4778
11.7539 | -94.4775
11.7542 | -94.4771
11.7545 | -94.4768
11.7548 | -94.4764
11.7551 | -94.4761
11.7554 | -94.4749
11.7557 | -94.4738
11.756 | -94.4728
11.7563 | -94.4721
11.7567 | -94.4715
11.757 | -94.4711
11.7573 | -94.4705
11.7576 | -94.4699
11.7579 | -94.4699
11.7582 | -94.4698
11.7585 | -94.4693
11.7588 | -94.469
11.7591 | -94.4689
11.7594 | -94.4685
11.7598 | -94.468
11.7601 | -94.4679
11.7604 | -94.4678
11.7607 | -94.4675
11.761 | -94.4666
11.7613 | -94.4662
11.7616 | -94.4659
11.7619 | -94.4655
11.7622 | -94.4651
11.7625 | -94.465
11.7629 | -94.4647
11.7632 | -94.4646
11.7635 | -94.4641
11.7638 | -94.4631
11.7641 | -94.4625
11.7644 | -94.4619
11.7647 | -94.4614
11.765 | -94.4609
11.7653 | -94.4605
11.7657 | -94.4601
11.766 | -94.4597
11.7663 | -94.4585
11.7666 | -94.4575
11.7669 | -94.457
11.7672 | -94.4567
11.7675 | -94.4566
11.7678 | -94.456
11.7681 | -94.4549
11.7685 | -94.4545
11.7688 | -94.4538
11.7691 | -94.4532
11.7694 | -94.4527
11.7697 | -94.452
11.77 | -94.4515
11.7703 | -94.451
11.7706 | -94.4506
11.7709 | -94.4504
11.7712 | -94.4499
11.7716 | -94.4494
11.7719 | -94.449
11.7722 | -94.4488
11.7725 | -94.4484
11.7728 | -94.4479
11.7731 | -94.4475
11.7734 | -94.447
11.7737 | -94.4464
11.774 | -94.4462
11.7744 | -94.4459
11.7747 | -94.4455
11.775 | -94.4449
11.7753 | -94.4444
11.7756 | -94.4442
11.7759 | -94.4444
11.7762 | -94.4444
11.7765 | -94.444
11.7768 | -94.4436
11.7771 | -94.4432
11.7775 | -94.4427
11.7778 | -94.4424
11.7781 | -94.442
11.7784 | -94.442
11.7787 | -94.4422
11.779 | -94.4419
11.7793 | -94.4415
11.7796 | -94.441
11.7799 | -94.4402
11.7803 | -94.4393
11.7806 | -94.4385
11.7809 | -94.4379
11.7812 | -94.4372
11.7815 | -94.4368
11.7818 | -94.4366
11.7821 | -94.4369
11.7824 | -94.4365
11.7827 | -94.4359
11.783 | -94.4354
11.7834 | -94.4352
11.7837 | -94.435
11.784 | -94.4345
11.7843 | -94.434
11.7846 | -94.4339
11.7849 | -94.4337
11.7852 | -94.4332
11.7855 | -94.4327
11.7858 | -94.4319
11.7862 | -94.4313
11.7865 | -94.4312
11.7868 | -94.4318
11.7871 | -94.4322
11.7874 | -94.4317
11.7877 | -94.4314
11.788 | -94.4309
11.7883 | -94.4306
11.7886 | -94.4303
11.7889 | -94.4302
11.7893 | -94.4296
11.7896 | -94.4287
11.7899 | -94.4278
11.7902 | -94.4273
11.7905 | -94.4266
11.7908 | -94.4254
11.7911 | -94.4243
11.7914 | -94.4235
11.7917 | -94.423
11.7921 | -94.423
11.7924 | -94.4225
11.7927 | -94.422
11.793 | -94.4219
11.7933 | -94.4214
11.7936 | -94.421
11.7939 | -94.4208
11.7942 | -94.4203
11.7945 | -94.4199
11.7948 | -94.4191
11.7952 | -94.4185
11.7955 | -94.4176
11.7958 | -94.4167
11.7961 | -94.4162
11.7964 | -94.4158
11.7967 | -94.4155
11.797 | -94.4153
11.7973 | -94.4153
11.7976 | -94.4155
11.798 | -94.4152
11.7983 | -94.4146
11.7986 | -94.4139
11.7989 | -94.4132
11.7992 | -94.4128
11.7995 | -94.4125
11.7998 | -94.4101
11.8001 | -94.4093
11.8004 | -94.4088
11.8007 | -94.4083
11.8011 | -94.408
11.8014 | -94.4073
11.8017 | -94.4063
11.802 | -94.4057
11.8023 | -94.4052
11.8026 | -94.405
11.8029 | -94.4053
11.8032 | -94.4051
11.8035 | -94.4046
11.8039 | -94.4041
11.8042 | -94.4041
11.8045 | -94.4036
11.8048 | -94.4032
11.8051 | -94.403
11.8054 | -94.4025
11.8057 | -94.402
11.806 | -94.4023
11.8063 | -94.402
11.8066 | -94.4015
11.807 | -94.401
11.8073 | -94.4006
11.8076 | -94.4002
11.8079 | -94.3992
11.8082 | -94.3984
11.8085 | -94.3981
11.8088 | -94.3979
11.8091 | -94.3973
11.8094 | -94.3964
11.8098 | -94.3958
11.8101 | -94.3956
11.8104 | -94.394
11.8107 | -94.3924
11.811 | -94.3917
11.8113 | -94.3911
11.8116 | -94.3901
11.8119 | -94.3895
11.8122 | -94.3894
11.8125 | -94.3892
11.8129 | -94.3889
11.8132 | -94.3885
11.8135 | -94.3878
11.8138 | -94.3874
11.8141 | -94.3872
11.8144 | -94.3873
11.8147 | -94.3874
11.815 | -94.3871
11.8153 | -94.3866
11.8157 | -94.386
11.816 | -94.3853
11.8163 | -94.3849
11.8166 | -94.3844
11.8169 | -94.3836
11.8172 | -94.3832
11.8175 | -94.3827
11.8178 | -94.3823
11.8181 | -94.382
11.8184 | -94.3817
11.8188 | -94.3814
11.8191 | -94.381
11.8194 | -94.3805
11.8197 | -94.3801
11.82 | -94.3798
11.8203 | -94.3793
11.8206 | -94.3789
11.8209 | -94.379
11.8212 | -94.3787
11.8216 | -94.3781
11.8219 | -94.3769
11.8222 | -94.3763
11.8225 | -94.3761
11.8228 | -94.3757
11.8231 | -94.3751
11.8234 | -94.3746
11.8237 | -94.3741
11.824 | -94.3735
11.8243 | -94.373
11.8247 | -94.3723
11.825 | -94.3716
11.8253 | -94.3705
11.8256 | -94.3698
11.8259 | -94.3689
11.8262 | -94.368
11.8265 | -94.3675
11.8268 | -94.3672
11.8271 | -94.3668
11.8275 | -94.3665
11.8278 | -94.3659
11.8281 | -94.3653
11.8284 | -94.3645
11.8287 | -94.364
11.829 | -94.3636
11.8293 | -94.363
11.8296 | -94.3624
11.8299 | -94.3622
11.8302 | -94.3627
11.8306 | -94.3625
11.8309 | -94.3622
11.8312 | -94.3626
11.8315 | -94.3623
11.8318 | -94.3618
11.8321 | -94.3616
11.8324 | -94.3612
11.8327 | -94.3606
11.833 | -94.3601
11.8334 | -94.3593
11.8337 | -94.3586
11.834 | -94.3581
11.8343 | -94.3577
11.8346 | -94.357
11.8349 | -94.3564
11.8352 | -94.3558
11.8355 | -94.3553
11.8358 | -94.3551
11.8361 | -94.3547
11.8365 | -94.3543
11.8368 | -94.3536
11.8371 | -94.353
11.8374 | -94.3526
11.8377 | -94.3523
11.838 | -94.3514
11.8383 | -94.3508
11.8386 | -94.3502
11.8389 | -94.3497
11.8393 | -94.3492
11.8396 | -94.3486
11.8399 | -94.3482
11.8402 | -94.3477
11.8405 | -94.3472
11.8408 | -94.3468
11.8411 | -94.3465
11.8414 | -94.3461
11.8417 | -94.3453
11.842 | -94.3447
11.8424 | -94.3441
11.8427 | -94.3435
11.843 | -94.3428
11.8433 | -94.3423
11.8436 | -94.3419
11.8439 | -94.3415
11.8442 | -94.3413
11.8445 | -94.341
11.8448 | -94.3404
11.8452 | -94.3398
11.8455 | -94.3391
11.8458 | -94.3388
11.8461 | -94.3385
11.8464 | -94.338
11.8467 | -94.3375
11.847 | -94.3367
11.8473 | -94.3363
11.8476 | -94.3357
11.8479 | -94.3353
11.8483 | -94.3347
11.8486 | -94.3341
11.8489 | -94.3336
11.8492 | -94.333
11.8495 | -94.3322
11.8498 | -94.3311
11.8501 | -94.3299
11.8504 | -94.3295
11.8507 | -94.3293
11.8511 | -94.3292
11.8514 | -94.329
11.8517 | -94.3285
11.852 | -94.3275
11.8523 | -94.327
11.8526 | -94.3265
11.8529 | -94.3258
11.8532 | -94.3254
11.8535 | -94.3251
11.8538 | -94.3243
11.8542 | -94.3236
11.8545 | -94.3228
11.8548 | -94.3222
11.8551 | -94.3218
11.8554 | -94.3213
11.8557 | -94.3208
11.856 | -94.32
11.8563 | -94.3196
11.8566 | -94.319
11.857 | -94.3183
11.8573 | -94.3178
11.8576 | -94.3172
11.8579 | -94.3166
11.8582 | -94.3161
11.8585 | -94.3148
11.8588 | -94.3133
11.8591 | -94.3126
11.8594 | -94.3124
11.8597 | -94.3121
11.8601 | -94.3112
11.8604 | -94.3106
11.8607 | -94.3101
11.861 | -94.3094
11.8613 | -94.309
11.8616 | -94.3087
11.8619 | -94.3084
11.8622 | -94.3075
11.8625 | -94.3069
11.8629 | -94.306
11.8632 | -94.3055
11.8635 | -94.3051
11.8638 | -94.3045
11.8641 | -94.3039
11.8644 | -94.3033
11.8647 | -94.3028
11.865 | -94.3024
11.8653 | -94.302
11.8656 | -94.3016
11.866 | -94.3014
11.8663 | -94.301
11.8666 | -94.3003
11.8669 | -94.2993
11.8672 | -94.2986
11.8675 | -94.2982
11.8678 | -94.2979
11.8681 | -94.2974
11.8684 | -94.297
11.8688 | -94.2965
11.8691 | -94.2958
11.8694 | -94.2952
11.8697 | -94.2947
11.87 | -94.2942
11.8703 | -94.294
11.8706 | -94.2938
11.8709 | -94.2932
11.8712 | -94.2927
11.8715 | -94.2925
11.8719 | -94.2919
11.8722 | -94.2917
11.8725 | -94.2909
11.8728 | -94.2903
11.8731 | -94.2899
11.8734 | -94.2896
11.8737 | -94.2893
11.874 | -94.2891
11.8743 | -94.2886
11.8747 | -94.2875
11.875 | -94.2867
11.8753 | -94.2862
11.8756 | -94.2859
11.8759 | -94.2851
11.8762 | -94.2842
11.8765 | -94.2829
11.8768 | -94.2824
11.8771 | -94.2819
11.8774 | -94.2811
11.8778 | -94.2805
11.8781 | -94.2805
11.8784 | -94.2801
11.8787 | -94.2795
11.879 | -94.2791
11.8793 | -94.2775
11.8796 | -94.276
11.8799 | -94.2756
11.8802 | -94.2752
11.8806 | -94.2738
11.8809 | -94.2726
11.8812 | -94.2726
11.8815 | -94.2721
11.8818 | -94.2718
11.8821 | -94.2716
11.8824 | -94.2711
11.8827 | -94.2707
11.883 | -94.2704
11.8834 | -94.2699
11.8837 | -94.2695
11.884 | -94.2688
11.8843 | -94.2685
11.8846 | -94.2675
11.8849 | -94.266
11.8852 | -94.265
11.8855 | -94.2646
11.8858 | -94.2639
11.8861 | -94.2633
11.8865 | -94.2625
11.8868 | -94.2618
11.8871 | -94.2613
11.8874 | -94.2603
11.8877 | -94.2593
11.888 | -94.2587
11.8883 | -94.2573
11.8886 | -94.2565
11.8889 | -94.2562
11.8893 | -94.2561
11.8896 | -94.255
11.8899 | -94.2541
11.8902 | -94.2536
11.8905 | -94.2536
11.8908 | -94.2536
11.8911 | -94.2529
11.8914 | -94.2524
11.8917 | -94.2519
11.892 | -94.2506
11.8924 | -94.2491
11.8927 | -94.2482
11.893 | -94.2477
11.8933 | -94.247
11.8936 | -94.246
11.8939 | -94.2448
11.8942 | -94.244
11.8945 | -94.2436
11.8948 | -94.243
11.8951 | -94.2426
11.8955 | -94.2426
11.8958 | -94.2419
11.8961 | -94.2403
11.8964 | -94.2393
11.8967 | -94.238
11.897 | -94.2372
11.8973 | -94.2366
11.8976 | -94.2351
11.8979 | -94.2335
11.8983 | -94.2326
11.8986 | -94.2324
11.8989 | -94.232
11.8992 | -94.2315
11.8995 | -94.2309
11.8998 | -94.2302
11.9001 | -94.2287
11.9004 | -94.2279
11.9007 | -94.2271
11.901 | -94.2266
11.9014 | -94.2261
11.9017 | -94.2254
11.902 | -94.2241
11.9023 | -94.2234
11.9026 | -94.223
11.9029 | -94.2226
11.9032 | -94.2219
11.9035 | -94.2209
11.9038 | -94.2197
11.9042 | -94.2191
11.9045 | -94.2185
11.9048 | -94.2179
11.9051 | -94.2173
11.9054 | -94.2163
11.9057 | -94.2155
11.906 | -94.2149
11.9063 | -94.2143
11.9066 | -94.2136
11.907 | -94.2128
11.9073 | -94.2111
11.9076 | -94.2095
11.9079 | -94.209
11.9082 | -94.2087
11.9085 | -94.2082
11.9088 | -94.2076
11.9091 | -94.2071
11.9094 | -94.2063
11.9097 | -94.2044
11.9101 | -94.2026
11.9104 | -94.2012
11.9107 | -94.2003
11.911 | -94.1988
11.9113 | -94.1981
11.9116 | -94.1975
11.9119 | -94.1971
11.9122 | -94.196
11.9125 | -94.1951
11.9129 | -94.1949
11.9132 | -94.1949
11.9135 | -94.1946
11.9138 | -94.1943
11.9141 | -94.1941
11.9144 | -94.1939
11.9147 | -94.1937
11.915 | -94.1938
11.9153 | -94.1934
11.9156 | -94.1923
11.916 | -94.1914
11.9163 | -94.1906
11.9166 | -94.1896
11.9169 | -94.1886
11.9172 | -94.1872
11.9175 | -94.1862
11.9178 | -94.1851
11.9181 | -94.1846
11.9184 | -94.1841
11.9188 | -94.1826
11.9191 | -94.1805
11.9194 | -94.1787
11.9197 | -94.1771
11.92 | -94.1759
11.9203 | -94.175
11.9206 | -94.1743
11.9209 | -94.1735
11.9212 | -94.1725
11.9215 | -94.1718
11.9219 | -94.1712
11.9222 | -94.1703
11.9225 | -94.1697
11.9228 | -94.1688
11.9231 | -94.1676
11.9234 | -94.1663
11.9237 | -94.165
11.924 | -94.1637
11.9243 | -94.1624
11.9247 | -94.1618
11.925 | -94.1612
11.9253 | -94.1603
11.9256 | -94.1603
11.9259 | -94.1595
11.9262 | -94.1587
11.9265 | -94.1579
11.9268 | -94.1568
11.9271 | -94.1561
11.9274 | -94.1559
11.9278 | -94.1555
11.9281 | -94.155
11.9284 | -94.1543
11.9287 | -94.1533
11.929 | -94.1525
11.9293 | -94.152
11.9296 | -94.152
11.9299 | -94.1514
11.9302 | -94.1506
11.9306 | -94.1499
11.9309 | -94.1496
11.9312 | -94.1491
11.9315 | -94.1484
11.9318 | -94.148
11.9321 | -94.1475
11.9324 | -94.1469
11.9327 | -94.1462
11.933 | -94.1454
11.9333 | -94.1447
11.9337 | -94.1439
11.934 | -94.1432
11.9343 | -94.1408
11.9346 | -94.1396
11.9349 | -94.139
11.9352 | -94.1387
11.9355 | -94.138
11.9358 | -94.1374
11.9361 | -94.1369
11.9365 | -94.1361
11.9368 | -94.1354
11.9371 | -94.1352
11.9374 | -94.1349
11.9377 | -94.1344
11.938 | -94.1336
11.9383 | -94.1327
11.9386 | -94.132
11.9389 | -94.1313
11.9392 | -94.1302
11.9396 | -94.1292
11.9399 | -94.1283
11.9402 | -94.1276
11.9405 | -94.1269
11.9408 | -94.126
11.9411 | -94.1249
11.9414 | -94.1242
11.9417 | -94.1236
11.942 | -94.1229
11.9424 | -94.1221
11.9427 | -94.1214
11.943 | -94.1206
11.9433 | -94.1199
11.9436 | -94.1195
11.9439 | -94.119
11.9442 | -94.1183
11.9445 | -94.1174
11.9448 | -94.1166
11.9451 | -94.1156
11.9455 | -94.115
11.9458 | -94.1143
11.9461 | -94.1137
11.9464 | -94.1129
11.9467 | -94.1122
11.947 | -94.1116
11.9473 | -94.1109
11.9476 | -94.1102
11.9479 | -94.1096
11.9483 | -94.1088
11.9486 | -94.1079
11.9489 | -94.1072
11.9492 | -94.1064
11.9495 | -94.1054
11.9498 | -94.1048
11.9501 | -94.1044
11.9504 | -94.1038
11.9507 | -94.1035
11.951 | -94.103
11.9514 | -94.1023
11.9517 | -94.1019
11.952 | -94.101
11.9523 | -94.0997
11.9526 | -94.0987
11.9529 | -94.0981
11.9532 | -94.0977
11.9535 | -94.0969
11.9538 | -94.0961
11.9542 | -94.0961
11.9545 | -94.0959
11.9548 | -94.0953
11.9551 | -94.0942
11.9554 | -94.0933
11.9557 | -94.0928
11.956 | -94.0922
11.9563 | -94.0914
11.9566 | -94.0908
11.9569 | -94.0902
11.9573 | -94.0897
11.9576 | -94.0891
11.9579 | -94.0884
11.9582 | -94.0878
11.9585 | -94.0873
11.9588 | -94.0868
11.9591 | -94.0862
11.9594 | -94.0855
11.9597 | -94.0847
11.9601 | -94.084
11.9604 | -94.0835
11.9607 | -94.0829
11.961 | -94.0825
11.9613 | -94.0819
11.9616 | -94.0815
11.9619 | -94.0812
11.9622 | -94.0809
11.9625 | -94.0801
11.9628 | -94.0793
11.9632 | -94.079
11.9635 | -94.0783
11.9638 | -94.0776
11.9641 | -94.077
11.9644 | -94.0763
11.9647 | -94.0756
11.965 | -94.0752
11.9653 | -94.0745
11.9656 | -94.0739
11.966 | -94.073
11.9663 | -94.0722
11.9666 | -94.0716
11.9669 | -94.0707
11.9672 | -94.0704
11.9675 | -94.0699
11.9678 | -94.0694
11.9681 | -94.069
11.9684 | -94.0685
11.9687 | -94.0681
11.9691 | -94.0673
11.9694 | -94.0666
11.9697 | -94.066
11.97 | -94.0654
11.9703 | -94.0647
11.9706 | -94.064
11.9709 | -94.0636
11.9712 | -94.0634
11.9715 | -94.0629
11.9719 | -94.0623
11.9722 | -94.0617
11.9725 | -94.0607
11.9728 | -94.06
11.9731 | -94.0595
11.9734 | -94.0589
11.9737 | -94.0583
11.974 | -94.0577
11.9743 | -94.0569
11.9746 | -94.0565
11.975 | -94.056
11.9753 | -94.0549
11.9756 | -94.0541
11.9759 | -94.0534
11.9762 | -94.0528
11.9765 | -94.0521
11.9768 | -94.0513
11.9771 | -94.0509
11.9774 | -94.0506
11.9778 | -94.0505
11.9781 | -94.05
11.9784 | -94.0495
11.9787 | -94.0489
11.979 | -94.0479
11.9793 | -94.0473
11.9796 | -94.0467
11.9799 | -94.046
11.9802 | -94.0454
11.9805 | -94.0449
11.9809 | -94.0443
11.9812 | -94.0437
11.9815 | -94.0429
11.9818 | -94.0422
11.9821 | -94.0417
11.9824 | -94.0409
11.9827 | -94.0402
11.983 | -94.0395
11.9833 | -94.0386
11.9837 | -94.0378
11.984 | -94.0374
11.9843 | -94.0371
11.9846 | -94.0369
11.9849 | -94.0364
11.9852 | -94.0358
11.9855 | -94.0354
11.9858 | -94.0351
11.9861 | -94.0346
11.9864 | -94.0337
11.9868 | -94.0328
11.9871 | -94.0323
11.9874 | -94.0321
11.9877 | -94.0318
11.988 | -94.0311
11.9883 | -94.0303
11.9886 | -94.0297
11.9889 | -94.0293
11.9892 | -94.0289
11.9896 | -94.0284
11.9899 | -94.0279
11.9902 | -94.0274
11.9905 | -94.0266
11.9908 | -94.026
11.9911 | -94.0255
11.9914 | -94.0252
11.9917 | -94.0246
11.992 | -94.0243
11.9923 | -94.0242
11.9927 | -94.0237
11.993 | -94.0231
11.9933 | -94.0226
11.9936 | -94.0219
11.9939 | -94.0212
11.9942 | -94.0205
11.9945 | -94.0199
11.9948 | -94.0195
11.9951 | -94.0189
11.9955 | -94.0182
11.9958 | -94.0178
11.9961 | -94.0171
11.9964 | -94.0164
11.9967 | -94.0156
11.997 | -94.0154
11.9973 | -94.0149
11.9976 | -94.0142
11.9979 | -94.0136
11.9982 | -94.0132
11.9986 | -94.0127
11.9989 | -94.0119
11.9992 | -94.0116
11.9995 | -94.0112
11.9998 | -94.011
12.0001 | -94.0109
12.0004 | -94.0104
12.0007 | -94.0094
12.001 | -94.0085
12.0014 | -94.0078
12.0017 | -94.0074
12.002 | -94.0069
12.0023 | -94.0063
12.0026 | -94.0059
12.0029 | -94.0053
12.0032 | -94.0048
12.0035 | -94.0042
12.0038 | -94.0031
12.0041 | -94.0027
12.0045 | -94.0023
12.0048 | -94.002
12.0051 | -94.0017
12.0054 | -94.001
12.0057 | -94.0002
12.006 | -93.9999
12.0063 | -93.9994
12.0066 | -93.999
12.0069 | -93.9983
12.0073 | -93.9977
12.0076 | -93.997
12.0079 | -93.9964
12.0082 | -93.9959
12.0085 | -93.9954
12.0088 | -93.9949
12.0091 | -93.9942
12.0094 | -93.9938
12.0097 | -93.9939
12.01 | -93.9933
12.0104 | -93.9925
12.0107 | -93.9919
12.011 | -93.9911
12.0113 | -93.9906
12.0116 | -93.9904
12.0119 | -93.9901
12.0122 | -93.9898
12.0125 | -93.9895
12.0128 | -93.9889
12.0132 | -93.9882
12.0135 | -93.9874
12.0138 | -93.9868
12.0141 | -93.9862
12.0144 | -93.9856
12.0147 | -93.9851
12.015 | -93.9847
12.0153 | -93.9838
12.0156 | -93.9828
12.0159 | -93.982
12.0163 | -93.9808
12.0166 | -93.9801
12.0169 | -93.9796
12.0172 | -93.9789
12.0175 | -93.9782
12.0178 | -93.9776
12.0181 | -93.9773
12.0184 | -93.977
12.0187 | -93.9765
12.0191 | -93.9756
12.0194 | -93.9751
12.0197 | -93.9747
12.02 | -93.974
12.0203 | -93.9733
12.0206 | -93.9727
12.0209 | -93.972
12.0212 | -93.9715
12.0215 | -93.9709
12.0218 | -93.9701
12.0222 | -93.9698
12.0225 | -93.9695
12.0228 | -93.9693
12.0231 | -93.9689
12.0234 | -93.9682
12.0237 | -93.968
12.024 | -93.9675
12.0243 | -93.9669
12.0246 | -93.9662
12.025 | -93.9657
12.0253 | -93.9652
12.0256 | -93.9645
12.0259 | -93.9639
12.0262 | -93.9633
12.0265 | -93.9628
12.0268 | -93.9622
12.0271 | -93.9614
12.0274 | -93.9608
12.0277 | -93.9608
12.0281 | -93.9603
12.0284 | -93.9598
12.0287 | -93.9592
12.029 | -93.9586
12.0293 | -93.9582
12.0296 | -93.9577
12.0299 | -93.9571
12.0302 | -93.9561
12.0305 | -93.9554
12.0309 | -93.9549
12.0312 | -93.9544
12.0315 | -93.9538
12.0318 | -93.9531
12.0321 | -93.952
12.0324 | -93.9515
12.0327 | -93.9511
12.033 | -93.9508
12.0333 | -93.9504
12.0336 | -93.9497
12.034 | -93.9487
12.0343 | -93.9482
12.0346 | -93.948
12.0349 | -93.9476
12.0352 | -93.9466
12.0355 | -93.9451
12.0358 | -93.9441
12.0361 | -93.9435
12.0364 | -93.9429
12.0368 | -93.9424
12.0371 | -93.9418
12.0374 | -93.9412
12.0377 | -93.9402
12.038 | -93.9397
12.0383 | -93.9392
12.0386 | -93.9387
12.0389 | -93.9383
12.0392 | -93.9377
12.0396 | -93.9375
12.0399 | -93.9371
12.0402 | -93.9364
12.0405 | -93.9361
12.0408 | -93.9356
12.0411 | -93.9348
12.0414 | -93.9343
12.0417 | -93.934
12.042 | -93.9338
12.0423 | -93.9337
12.0427 | -93.9333
12.043 | -93.9329
12.0433 | -93.9324
12.0436 | -93.932
12.0439 | -93.9314
12.0442 | -93.9305
12.0445 | -93.93
12.0448 | -93.9293
12.0451 | -93.9287
12.0455 | -93.9283
12.0458 | -93.9276
12.0461 | -93.9271
12.0464 | -93.9265
12.0467 | -93.926
12.047 | -93.9255
12.0473 | -93.9249
12.0476 | -93.9242
12.0479 | -93.9235
12.0482 | -93.9231
12.0486 | -93.9227
12.0489 | -93.9222
12.0492 | -93.9219
12.0495 | -93.9214
12.0498 | -93.9209
12.0501 | -93.9203
12.0504 | -93.9198
12.0507 | -93.9186
12.051 | -93.9182
12.0514 | -93.9179
12.0517 | -93.9176
12.052 | -93.917
12.0523 | -93.9162
12.0526 | -93.9157
12.0529 | -93.9152
12.0532 | -93.9147
12.0535 | -93.9143
12.0538 | -93.9137
12.0541 | -93.9129
12.0545 | -93.9124
12.0548 | -93.912
12.0551 | -93.9119
12.0554 | -93.9117
12.0557 | -93.9113
12.056 | -93.9111
12.0563 | -93.9108
12.0566 | -93.9102
12.0569 | -93.9094
12.0573 | -93.9088
12.0576 | -93.9083
12.0579 | -93.9077
12.0582 | -93.9068
12.0585 | -93.9062
12.0588 | -93.9055
12.0591 | -93.9051
12.0594 | -93.9048
12.0597 | -93.9047
12.06 | -93.904
12.0604 | -93.9034
12.0607 | -93.9033
12.061 | -93.9035
12.0613 | -93.9031
12.0616 | -93.9023
12.0619 | -93.9015
12.0622 | -93.9009
12.0625 | -93.9004
12.0628 | -93.8997
12.0632 | -93.8995
12.0635 | -93.8996
12.0638 | -93.8993
12.0641 | -93.899
12.0644 | -93.8988
12.0647 | -93.8982
12.065 | -93.8975
12.0653 | -93.897
12.0656 | -93.8966
12.0659 | -93.8963
12.0663 | -93.8961
12.0666 | -93.8959
12.0669 | -93.8955
12.0672 | -93.895
12.0675 | -93.8945
12.0678 | -93.894
12.0681 | -93.8937
12.0684 | -93.8933
12.0687 | -93.8931
12.0691 | -93.8926
12.0694 | -93.8921
12.0697 | -93.8915
12.07 | -93.8911
12.0703 | -93.8912
12.0706 | -93.8906
12.0709 | -93.8902
12.0712 | -93.8897
12.0715 | -93.8892
12.0718 | -93.8887
12.0722 | -93.8882
12.0725 | -93.8879
12.0728 | -93.8876
12.0731 | -93.887
12.0734 | -93.8864
12.0737 | -93.8859
12.074 | -93.8853
12.0743 | -93.8847
12.0746 | -93.8838
12.075 | -93.8829
12.0753 | -93.8821
12.0756 | -93.8815
12.0759 | -93.8813
12.0762 | -93.8805
12.0765 | -93.8795
12.0768 | -93.878
12.0771 | -93.8757
12.0774 | -93.8742
12.0777 | -93.8718
12.0781 | -93.8695
12.0784 | -93.8669
12.0787 | -93.8613
12.079 | -93.8583
12.0793 | -93.8576
12.0796 | -93.857
12.0799 | -93.8563
12.0802 | -93.8548
12.0805 | -93.8491
12.0809 | -93.8466
12.0812 | -93.8458
12.0815 | -93.8431
12.0818 | -93.8401
12.0821 | -93.839
12.0824 | -93.838
12.0827 | -93.8373
12.083 | -93.8368
12.0833 | -93.8362
12.0836 | -93.8356
12.084 | -93.8339
12.0843 | -93.8331
12.0846 | -93.8323
12.0849 | -93.8311
12.0852 | -93.8294
12.0855 | -93.8284
12.0858 | -93.8276
12.0861 | -93.8269
12.0864 | -93.8264
12.0868 | -93.8255
12.0871 | -93.8243
12.0874 | -93.8226
12.0877 | -93.8216
12.088 | -93.8207
12.0883 | -93.8199
12.0886 | -93.8188
12.0889 | -93.8177
12.0892 | -93.8172
12.0895 | -93.8166
12.0899 | -93.8162
12.0902 | -93.8154
12.0905 | -93.8147
12.0908 | -93.8142
12.0911 | -93.8141
12.0914 | -93.8139
12.0917 | -93.8134
12.092 | -93.8128
12.0923 | -93.8121
12.0927 | -93.8115
12.093 | -93.8111
12.0933 | -93.8105
12.0936 | -93.8096
12.0939 | -93.8089
12.0942 | -93.8086
12.0945 | -93.8081
12.0948 | -93.8072
12.0951 | -93.806
12.0954 | -93.8052
12.0958 | -93.8048
12.0961 | -93.8043
12.0964 | -93.8032
12.0967 | -93.8017
12.097 | -93.8006
12.0973 | -93.7999
12.0976 | -93.7991
12.0979 | -93.7985
12.0982 | -93.7979
12.0986 | -93.7973
12.0989 | -93.7967
12.0992 | -93.7962
12.0995 | -93.7959
12.0998 | -93.7953
12.1001 | -93.7943
12.1004 | -93.7935
12.1007 | -93.7931
12.101 | -93.7922
12.1013 | -93.7907
12.1017 | -93.7901
12.102 | -93.7895
12.1023 | -93.7886
12.1026 | -93.7879
12.1029 | -93.787
12.1032 | -93.7858
12.1035 | -93.7846
12.1038 | -93.7838
12.1041 | -93.7833
12.1045 | -93.7826
12.1048 | -93.7819
12.1051 | -93.7812
12.1054 | -93.7806
12.1057 | -93.78
12.106 | -93.7792
12.1063 | -93.7786
12.1066 | -93.778
12.1069 | -93.7776
12.1072 | -93.7769
12.1076 | -93.7761
12.1079 | -93.7754
12.1082 | -93.775
12.1085 | -93.7747
12.1088 | -93.7741
12.1091 | -93.7734
12.1094 | -93.7722
12.1097 | -93.7711
12.11 | -93.7704
12.1104 | -93.77
12.1107 | -93.7696
12.111 | -93.769
12.1113 | -93.7685
12.1116 | -93.7677
12.1119 | -93.767
12.1122 | -93.7663
12.1125 | -93.7657
12.1128 | -93.7651
12.1131 | -93.7644
12.1135 | -93.7639
12.1138 | -93.7633
12.1141 | -93.7628
12.1144 | -93.7623
12.1147 | -93.7618
12.115 | -93.7611
12.1153 | -93.7602
12.1156 | -93.7596
12.1159 | -93.7591
12.1163 | -93.7588
12.1166 | -93.7584
12.1169 | -93.7579
12.1172 | -93.7573
12.1175 | -93.7558
12.1178 | -93.755
12.1181 | -93.7542
12.1184 | -93.7538
12.1187 | -93.7528
12.119 | -93.7518
12.1194 | -93.7511
12.1197 | -93.7504
12.12 | -93.7497
12.1203 | -93.7487
12.1206 | -93.7478
12.1209 | -93.7475
12.1212 | -93.7472
12.1215 | -93.7467
12.1218 | -93.7461
12.1222 | -93.7456
12.1225 | -93.745
12.1228 | -93.7444
12.1231 | -93.7439
12.1234 | -93.7438
12.1237 | -93.7437
12.124 | -93.7432
12.1243 | -93.7424
12.1246 | -93.7416
12.1249 | -93.7408
12.1253 | -93.7403
12.1256 | -93.7401
12.1259 | -93.7395
12.1262 | -93.739
12.1265 | -93.7387
12.1268 | -93.7383
12.1271 | -93.7377
12.1274 | -93.7372
12.1277 | -93.7366
12.1281 | -93.7362
12.1284 | -93.7355
12.1287 | -93.7345
12.129 | -93.7337
12.1293 | -93.7332
12.1296 | -93.7325
12.1299 | -93.7319
12.1302 | -93.7311
12.1305 | -93.7305
12.1308 | -93.7301
12.1312 | -93.7295
12.1315 | -93.7291
12.1318 | -93.7284
12.1321 | -93.7274
12.1324 | -93.7263
12.1327 | -93.7257
12.133 | -93.7252
12.1333 | -93.7247
12.1336 | -93.7241
12.134 | -93.7236
12.1343 | -93.7229
12.1346 | -93.7223
12.1349 | -93.722
12.1352 | -93.7215
12.1355 | -93.7209
12.1358 | -93.7205
12.1361 | -93.7202
12.1364 | -93.7199
12.1367 | -93.7196
12.1371 | -93.7191
12.1374 | -93.7182
12.1377 | -93.7176
12.138 | -93.7174
12.1383 | -93.7169
12.1386 | -93.7162
12.1389 | -93.7158
12.1392 | -93.7153
12.1395 | -93.7145
12.1399 | -93.7138
12.1402 | -93.7133
12.1405 | -93.7128
12.1408 | -93.7122
12.1411 | -93.7116
12.1414 | -93.7108
12.1417 | -93.7103
12.142 | -93.7098
12.1423 | -93.7094
12.1426 | -93.7091
12.143 | -93.7086
12.1433 | -93.7077
12.1436 | -93.7071
12.1439 | -93.7069
12.1442 | -93.7065
12.1445 | -93.706
12.1448 | -93.7054
12.1451 | -93.7047
12.1454 | -93.7043
12.1458 | -93.7046
12.1461 | -93.7045
12.1464 | -93.7041
12.1467 | -93.7037
12.147 | -93.7031
12.1473 | -93.7025
12.1476 | -93.7016
12.1479 | -93.7009
12.1482 | -93.7004
12.1485 | -93.6996
12.1489 | -93.699
12.1492 | -93.6985
12.1495 | -93.6977
12.1498 | -93.6972
12.1501 | -93.6967
12.1504 | -93.6963
12.1507 | -93.6955
12.151 | -93.6948
12.1513 | -93.6942
12.1517 | -93.6936
12.152 | -93.6932
12.1523 | -93.6929
12.1526 | -93.6921
12.1529 | -93.6915
12.1532 | -93.6906
12.1535 | -93.6898
12.1538 | -93.6894
12.1541 | -93.6889
12.1544 | -93.6885
12.1548 | -93.6881
12.1551 | -93.6879
12.1554 | -93.6878
12.1557 | -93.6872
12.156 | -93.6865
12.1563 | -93.686
12.1566 | -93.6855
12.1569 | -93.6847
12.1572 | -93.6839
12.1576 | -93.6832
12.1579 | -93.6831
12.1582 | -93.6827
12.1585 | -93.6824
12.1588 | -93.6819
12.1591 | -93.6815
12.1594 | -93.681
12.1597 | -93.6802
12.16 | -93.6794
12.1603 | -93.6787
12.1607 | -93.6778
12.161 | -93.677
12.1613 | -93.6763
12.1616 | -93.6755
12.1619 | -93.6749
12.1622 | -93.6743
12.1625 | -93.6737
12.1628 | -93.6729
12.1631 | -93.6722
12.1635 | -93.6716
12.1638 | -93.6711
12.1641 | -93.6706
12.1644 | -93.6695
12.1647 | -93.6688
12.165 | -93.6682
12.1653 | -93.6676
12.1656 | -93.6671
12.1659 | -93.6661
12.1662 | -93.665
12.1666 | -93.6644
12.1669 | -93.664
12.1672 | -93.6638
12.1675 | -93.6632
12.1678 | -93.6626
12.1681 | -93.6624
12.1684 | -93.6618
12.1687 | -93.661
12.169 | -93.6603
12.1694 | -93.6596
12.1697 | -93.6589
12.17 | -93.6582
12.1703 | -93.6576
12.1706 | -93.657
12.1709 | -93.6565
12.1712 | -93.6557
12.1715 | -93.6551
12.1718 | -93.6545
12.1722 | -93.654
12.1725 | -93.6533
12.1728 | -93.6529
12.1731 | -93.6524
12.1734 | -93.6518
12.1737 | -93.6511
12.174 | -93.6508
12.1743 | -93.6504
12.1746 | -93.6496
12.1749 | -93.6478
12.1753 | -93.646
12.1756 | -93.6451
12.1759 | -93.6443
12.1762 | -93.6437
12.1765 | -93.6432
12.1768 | -93.6428
12.1771 | -93.6422
12.1774 | -93.6417
12.1777 | -93.6411
12.1781 | -93.6404
12.1784 | -93.6397
12.1787 | -93.6392
12.179 | -93.6387
12.1793 | -93.638
12.1796 | -93.6372
12.1799 | -93.6367
12.1802 | -93.6362
12.1805 | -93.6357
12.1808 | -93.6354
12.1812 | -93.6349
12.1815 | -93.6343
12.1818 | -93.6337
12.1821 | -93.6333
12.1824 | -93.633
12.1827 | -93.6325
12.183 | -93.632
12.1833 | -93.6316
12.1836 | -93.631
12.1839 | -93.6308
12.1843 | -93.6304
12.1846 | -93.63
12.1849 | -93.6295
12.1852 | -93.629
12.1855 | -93.6282
12.1858 | -93.6273
12.1861 | -93.6266
12.1864 | -93.6261
12.1867 | -93.6257
12.1871 | -93.6255
12.1874 | -93.6253
12.1877 | -93.6248
12.188 | -93.6243
12.1883 | -93.6237
12.1886 | -93.6232
12.1889 | -93.623
12.1892 | -93.6226
12.1895 | -93.622
12.1898 | -93.6215
12.1902 | -93.6208
12.1905 | -93.6203
12.1908 | -93.6198
12.1911 | -93.6193
12.1914 | -93.6183
12.1917 | -93.6174
12.192 | -93.6168
12.1923 | -93.6161
12.1926 | -93.6157
12.193 | -93.6152
12.1933 | -93.6146
12.1936 | -93.6142
12.1939 | -93.6139
12.1942 | -93.6136
12.1945 | -93.6131
12.1948 | -93.6125
12.1951 | -93.6117
12.1954 | -93.611
12.1958 | -93.6104
12.1961 | -93.6101
12.1964 | -93.61
12.1967 | -93.6097
12.197 | -93.6093
12.1973 | -93.6091
12.1976 | -93.6087
12.1979 | -93.6084
12.1982 | -93.6087
12.1985 | -93.6088
12.1989 | -93.6086
12.1992 | -93.6083
12.1995 | -93.6075
12.1998 | -93.6061
12.2001 | -93.6053
12.2004 | -93.605
12.2007 | -93.6042
12.201 | -93.6036
12.2013 | -93.6031
12.2017 | -93.6024
12.202 | -93.6019
12.2023 | -93.6013
12.2026 | -93.6007
12.2029 | -93.5999
12.2032 | -93.5996
12.2035 | -93.5989
12.2038 | -93.5983
12.2041 | -93.5975
12.2044 | -93.597
12.2048 | -93.5968
12.2051 | -93.5964
12.2054 | -93.5959
12.2057 | -93.5948
12.206 | -93.5934
12.2063 | -93.5925
12.2066 | -93.592
12.2069 | -93.5912
12.2072 | -93.5904
12.2076 | -93.5899
12.2079 | -93.5896
12.2082 | -93.589
12.2085 | -93.5886
12.2088 | -93.588
12.2091 | -93.5875
12.2094 | -93.5868
12.2097 | -93.5862
12.21 | -93.5858
12.2103 | -93.5853
12.2107 | -93.5847
12.211 | -93.5837
12.2113 | -93.5828
12.2116 | -93.5822
12.2119 | -93.582
12.2122 | -93.5816
12.2125 | -93.5811
12.2128 | -93.5805
12.2131 | -93.5803
12.2135 | -93.58
12.2138 | -93.5795
12.2141 | -93.5789
12.2144 | -93.5782
12.2147 | -93.5777
12.215 | -93.5769
12.2153 | -93.5767
12.2156 | -93.5762
12.2159 | -93.5756
12.2162 | -93.575
12.2166 | -93.5745
12.2169 | -93.5738
12.2172 | -93.5733
12.2175 | -93.5728
12.2178 | -93.5722
12.2181 | -93.5713
12.2184 | -93.5708
12.2187 | -93.5704
12.219 | -93.57
12.2194 | -93.5696
12.2197 | -93.5692
12.22 | -93.5682
12.2203 | -93.5678
12.2206 | -93.5675
12.2209 | -93.5668
12.2212 | -93.5663
12.2215 | -93.5658
12.2218 | -93.5653
12.2221 | -93.5648
12.2225 | -93.5642
12.2228 | -93.564
12.2231 | -93.5635
12.2234 | -93.5628
12.2237 | -93.5623
12.224 | -93.5619
12.2243 | -93.5617
12.2246 | -93.5612
12.2249 | -93.5607
12.2253 | -93.5605
12.2256 | -93.5604
12.2259 | -93.5602
12.2262 | -93.5597
12.2265 | -93.5589
12.2268 | -93.5578
12.2271 | -93.5569
12.2274 | -93.5564
12.2277 | -93.5558
12.228 | -93.5553
12.2284 | -93.5547
12.2287 | -93.5543
12.229 | -93.5537
12.2293 | -93.5531
12.2296 | -93.5527
12.2299 | -93.5523
12.2302 | -93.5519
12.2305 | -93.5516
12.2308 | -93.5515
12.2312 | -93.5512
12.2315 | -93.5506
12.2318 | -93.5502
12.2321 | -93.5499
12.2324 | -93.5496
12.2327 | -93.549
12.233 | -93.5489
12.2333 | -93.5487
12.2336 | -93.5483
12.2339 | -93.548
12.2343 | -93.5475
12.2346 | -93.5473
12.2349 | -93.5472
12.2352 | -93.5469
12.2355 | -93.5463
12.2358 | -93.5458
12.2361 | -93.5451
12.2364 | -93.5439
12.2367 | -93.5433
12.2371 | -93.5426
12.2374 | -93.5423
12.2377 | -93.5418
12.238 | -93.5412
12.2383 | -93.5405
12.2386 | -93.5393
12.2389 | -93.5388
12.2392 | -93.5383
12.2395 | -93.5378
12.2398 | -93.5373
12.2402 | -93.5368
12.2405 | -93.5366
12.2408 | -93.5363
12.2411 | -93.5359
12.2414 | -93.5354
12.2417 | -93.5353
12.242 | -93.535
12.2423 | -93.5342
12.2426 | -93.5329
12.243 | -93.5325
12.2433 | -93.5322
12.2436 | -93.5317
12.2439 | -93.5313
12.2442 | -93.5311
12.2445 | -93.5305
12.2448 | -93.53
12.2451 | -93.5295
12.2454 | -93.5288
12.2457 | -93.528
12.2461 | -93.5274
12.2464 | -93.5268
12.2467 | -93.5263
12.247 | -93.5259
12.2473 | -93.5252
12.2476 | -93.5245
12.2479 | -93.5245
12.2482 | -93.5236
12.2485 | -93.5231
12.2489 | -93.5228
12.2492 | -93.5226
12.2495 | -93.5221
12.2498 | -93.5216
12.2501 | -93.521
12.2504 | -93.5205
12.2507 | -93.52
12.251 | -93.5199
12.2513 | -93.5197
12.2516 | -93.5194
12.252 | -93.5191
12.2523 | -93.5187
12.2526 | -93.518
12.2529 | -93.5175
12.2532 | -93.517
12.2535 | -93.5166
12.2538 | -93.5163
12.2541 | -93.5162
12.2544 | -93.5158
12.2548 | -93.5151
12.2551 | -93.5147
12.2554 | -93.5143
12.2557 | -93.5138
12.256 | -93.5134
12.2563 | -93.513
12.2566 | -93.5125
12.2569 | -93.5122
12.2572 | -93.5118
12.2575 | -93.5115
12.2579 | -93.5114
12.2582 | -93.5114
12.2585 | -93.511
12.2588 | -93.5106
12.2591 | -93.5103
12.2594 | -93.5098
12.2597 | -93.5094
12.26 | -93.5085
12.2603 | -93.5076
12.2607 | -93.5072
12.261 | -93.5069
12.2613 | -93.5068
12.2616 | -93.5064
12.2619 | -93.506
12.2622 | -93.5055
12.2625 | -93.5048
12.2628 | -93.5045
12.2631 | -93.504
12.2634 | -93.5037
12.2638 | -93.5032
12.2641 | -93.5026
12.2644 | -93.5016
12.2647 | -93.5007
12.265 | -93.5004
12.2653 | -93.4999
12.2656 | -93.4994
12.2659 | -93.4992
12.2662 | -93.4988
12.2666 | -93.4984
12.2669 | -93.498
12.2672 | -93.4976
12.2675 | -93.4972
12.2678 | -93.4968
12.2681 | -93.4963
12.2684 | -93.4958
12.2687 | -93.4952
12.269 | -93.4946
12.2693 | -93.4942
12.2697 | -93.494
12.27 | -93.4936
12.2703 | -93.4934
12.2706 | -93.493
12.2709 | -93.4926
12.2712 | -93.4921
12.2715 | -93.4916
12.2718 | -93.4913
12.2721 | -93.4908
12.2725 | -93.4903
12.2728 | -93.4898
12.2731 | -93.4895
12.2734 | -93.4889
12.2737 | -93.488
12.274 | -93.4874
12.2743 | -93.487
12.2746 | -93.4869
12.2749 | -93.4865
12.2752 | -93.4858
12.2756 | -93.4854
12.2759 | -93.4852
12.2762 | -93.4849
12.2765 | -93.4847
12.2768 | -93.4845
12.2771 | -93.4842
12.2774 | -93.4839
12.2777 | -93.4834
12.278 | -93.4831
12.2784 | -93.4826
12.2787 | -93.4823
12.279 | -93.4823
12.2793 | -93.4823
12.2796 | -93.4819
12.2799 | -93.481
12.2802 | -93.4806
12.2805 | -93.4803
12.2808 | -93.48
12.2811 | -93.4797
12.2815 | -93.479
12.2818 | -93.4783
12.2821 | -93.4782
12.2824 | -93.4781
12.2827 | -93.4778
12.283 | -93.4777
12.2833 | -93.477
12.2836 | -93.4758
12.2839 | -93.4754
12.2843 | -93.4751
12.2846 | -93.4748
12.2849 | -93.4744
12.2852 | -93.4741
12.2855 | -93.4737
12.2858 | -93.4732
12.2861 | -93.4727
12.2864 | -93.4723
12.2867 | -93.4723
12.287 | -93.4721
12.2874 | -93.4717
12.2877 | -93.4715
12.288 | -93.4712
12.2883 | -93.4708
12.2886 | -93.4706
12.2889 | -93.4701
12.2892 | -93.4699
12.2895 | -93.4695
12.2898 | -93.4688
12.2902 | -93.4681
12.2905 | -93.4675
12.2908 | -93.4672
12.2911 | -93.4671
12.2914 | -93.4668
12.2917 | -93.4664
12.292 | -93.4661
12.2923 | -93.4657
12.2926 | -93.4655
12.2929 | -93.4653
12.2933 | -93.4653
12.2936 | -93.4649
12.2939 | -93.4648
12.2942 | -93.4641
12.2945 | -93.4636
12.2948 | -93.4632
12.2951 | -93.4628
12.2954 | -93.4624
12.2957 | -93.4619
12.2961 | -93.4615
12.2964 | -93.4612
12.2967 | -93.4609
12.297 | -93.4605
12.2973 | -93.4601
12.2976 | -93.4595
12.2979 | -93.4592
12.2982 | -93.4593
12.2985 | -93.4591
12.2988 | -93.4589
12.2992 | -93.4588
12.2995 | -93.4585
12.2998 | -93.4581
12.3001 | -93.4578
12.3004 | -93.4574
12.3007 | -93.457
12.301 | -93.4568
12.3013 | -93.4565
12.3016 | -93.4567
12.302 | -93.4564
12.3023 | -93.4562
12.3026 | -93.456
12.3029 | -93.4558
12.3032 | -93.4556
12.3035 | -93.4553
12.3038 | -93.455
12.3041 | -93.4549
12.3044 | -93.4549
12.3047 | -93.4547
12.3051 | -93.4542
12.3054 | -93.4537
12.3057 | -93.4535
12.306 | -93.4533
12.3063 | -93.4527
12.3066 | -93.4524
12.3069 | -93.4521
12.3072 | -93.4517
12.3075 | -93.451
12.3079 | -93.4503
12.3082 | -93.4497
12.3085 | -93.4495
12.3088 | -93.4496
12.3091 | -93.4495
12.3094 | -93.4495
12.3097 | -93.4492
12.31 | -93.4487
12.3103 | -93.4478
12.3107 | -93.4477
12.311 | -93.4474
12.3113 | -93.4471
12.3116 | -93.4468
12.3119 | -93.4466
12.3122 | -93.4464
12.3125 | -93.4466
12.3128 | -93.4469
12.3131 | -93.4471
12.3134 | -93.4466
12.3138 | -93.4462
12.3141 | -93.4461
12.3144 | -93.4459
12.3147 | -93.4456
12.315 | -93.4452
12.3153 | -93.4448
12.3156 | -93.4453
12.3159 | -93.446
12.3162 | -93.4458
12.3165 | -93.4453
12.3169 | -93.445
12.3172 | -93.4446
12.3175 | -93.4443
12.3178 | -93.4441
12.3181 | -93.444
12.3184 | -93.4438
12.3187 | -93.4435
12.319 | -93.4433
12.3193 | -93.443
12.3197 | -93.4427
12.32 | -93.4423
12.3203 | -93.442
12.3206 | -93.4416
12.3209 | -93.4411
12.3212 | -93.4403
12.3215 | -93.4396
12.3218 | -93.4391
12.3221 | -93.4392
12.3224 | -93.4399
12.3228 | -93.4396
12.3231 | -93.439
12.3234 | -93.4386
12.3237 | -93.4381
12.324 | -93.4378
12.3243 | -93.4377
12.3246 | -93.4375
12.3249 | -93.4379
12.3252 | -93.4377
12.3256 | -93.4374
12.3259 | -93.4372
12.3262 | -93.4369
12.3265 | -93.4359
12.3268 | -93.4353
12.3271 | -93.4353
12.3274 | -93.4353
12.3277 | -93.4351
12.328 | -93.4347
12.3284 | -93.4343
12.3287 | -93.4338
12.329 | -93.4334
12.3293 | -93.4332
12.3296 | -93.4327
12.3299 | -93.4322
12.3302 | -93.4319
12.3305 | -93.4317
12.3308 | -93.4313
12.3311 | -93.4312
12.3315 | -93.4309
12.3318 | -93.4305
12.3321 | -93.4303
12.3324 | -93.4298
12.3327 | -93.4294
12.333 | -93.4288
12.3333 | -93.4284
12.3336 | -93.4282
12.3339 | -93.4278
12.3343 | -93.4273
12.3346 | -93.427
12.3349 | -93.427
12.3352 | -93.4269
12.3355 | -93.4267
12.3358 | -93.4268
12.3361 | -93.4264
12.3364 | -93.4261
12.3367 | -93.4257
12.337 | -93.4252
12.3374 | -93.4246
12.3377 | -93.4247
12.338 | -93.4244
12.3383 | -93.424
12.3386 | -93.4236
12.3389 | -93.4235
12.3392 | -93.4233
12.3395 | -93.4232
12.3398 | -93.4229
12.3402 | -93.4226
12.3405 | -93.4224
12.3408 | -93.4221
12.3411 | -93.4218
12.3414 | -93.4212
12.3417 | -93.4207
12.342 | -93.4205
12.3423 | -93.4202
12.3426 | -93.42
12.3429 | -93.4199
12.3433 | -93.4195
12.3436 | -93.4193
12.3439 | -93.4193
12.3442 | -93.4192
12.3445 | -93.4188
12.3448 | -93.4189
12.3451 | -93.4185
12.3454 | -93.4182
12.3457 | -93.4177
12.3461 | -93.4171
12.3464 | -93.4173
12.3467 | -93.4169
12.347 | -93.4166
12.3473 | -93.4167
12.3476 | -93.4164
12.3479 | -93.4161
12.3482 | -93.4158
12.3485 | -93.4154
12.3488 | -93.4151
12.3492 | -93.4143
12.3495 | -93.414
12.3498 | -93.414
12.3501 | -93.4139
12.3504 | -93.4134
12.3507 | -93.4129
12.351 | -93.4127
12.3513 | -93.4125
12.3516 | -93.4124
12.352 | -93.4119
12.3523 | -93.4113
12.3526 | -93.411
12.3529 | -93.4106
12.3532 | -93.4105
12.3535 | -93.4103
12.3538 | -93.41
12.3541 | -93.4096
12.3544 | -93.4088
12.3547 | -93.4082
12.3551 | -93.4078
12.3554 | -93.4069
12.3557 | -93.4067
12.356 | -93.4065
12.3563 | -93.4063
12.3566 | -93.4063
12.3569 | -93.4061
12.3572 | -93.4065
12.3575 | -93.4065
12.3579 | -93.4063
12.3582 | -93.4066
12.3585 | -93.4068
12.3588 | -93.4064
12.3591 | -93.4061
12.3594 | -93.4059
12.3597 | -93.4056
12.36 | -93.4054
12.3603 | -93.4053
12.3606 | -93.4056
12.361 | -93.4055
12.3613 | -93.4054
12.3616 | -93.4055
12.3619 | -93.4054
12.3622 | -93.4052
12.3625 | -93.4051
12.3628 | -93.4048
12.3631 | -93.4044
12.3634 | -93.4039
12.3638 | -93.404