## CALIBRATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 71 |  | CTBTO
Transducer Block | 02-90C/P | 68 |  | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
Hood | 02-92B/P | 57
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE REFERENCE
Atmospheric pressure | 1013.8 | hPa |  | PR 000 150_-_CAL
Humidity | 19.3 | %/Hr |  | Measurement device
Internal Temperature of Calibration Bench | 22.62 | °C |  | IS CALIBRATOR
AUX Sensor pressure | 1015 | hPa |  | Sensor SoH at test start
AUX Sensor Temperature | 22.81 | °C |  | OK
SENSITIVITY AT 1 Hz
Output Parameters | Theoretical Value | Measured Value @1Hz | Margin (±) | Unit | Result
Sensitivity of Pressure Derivative output | 2 | 1.99 | 0.04 | mV/Pa/sec | OK | Mode d'emploi pour créer un constat de vérification
Sensitivity of Pressure output | 20 | 19.95 | 0.4 | mV/Pa | OK | Cliquer sur "DEMARRER LA SAISIE"
Sensitivity of Calibration coil | 6 | 5.96 | 0.13 | Pa/V | OK | Suivre les instructions affichées
Une fois la partie SENSOR IDENTIFICATION remplie, renseigner la partie ENVIRONNEMENTAL PARAMETERS (les données environnementales)
TRANSFER FUNCTION | [5 empty] | Renseigner la partie SENSITIVITY AT 1Hz :les sensibilités en voie de dérivée de Pression (Pressure Derivative), en voie de pression (Pressure) et de Calibration (calibration Coil)
Electrical |  | Acoustical |  |  |  | Une sauvegarde temporaire est toujours possible si le fichier doit être modifié (Fichier "enregistrer")
Vérifier toutes les données saisie notamment la partie CUSTOMER ("Client") et la partie VISA (Noms des personnes et date d'ajustage)
Cliquer sur "ENREGISTREMENT DEFINITIF" (Attention le fichier sera enregistré et verrouillé dans le dossier courant, il faut retirer manuellement la protection)
SENSOR SELF NOISE |  |  |  | Decision : considering the specifications given by the manufacturer, this sensor is
CONFORM
VISA
|  |  |  | Adjusted by | A.HUE
|  |  |  | on | 2016-03-14
|  |  |  | Approved by | A.HUE
|  |  |  | Qualification | Optronic Engineer
Self Noise at 1 HZ | 0.00010328 | Pa/√Hz |  | on | 2016-03-17 | -79.72 | (dBPa/√Hz)
| < 0,13 | mPa/√Hz
OK
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which define the verification report. It can be used to demonstrate the link between the measure equipment and the standard national or international equipment, provided it meets the recommendations of the documentation booklet X 07-015
[8 empty] | Client | Adresse Client
[8 empty] | CEA | C.E.A. Dam. Bruyères Domaine du Grand Rué 91297 ARPAJON CEDEX FRANCE
[8 empty] | CTBTO | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
[8 empty] | IPGP | Université Paris Diderot Paris 7 Bat Lamarck UFR STEP 35 rue H.Brion 6ème étage 75205 Paris CEDEX 13
[8 empty] | SANDIA | C/O SANDIATIONAL LABS ALBUQUERQUE, NM 87123 UNITED STATES
[8 empty] | ISLA | Haiwaï Institute of Geophysics & Planetology 73-984 Makako Bay Drive KAILUA-KONA HAWAÏ
ENREGISTREMENT DEFINITIF
DEMARRER LA SAISIE

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.6558 | 9.4246
0.097656 | -34.0665 | 7.2697
0.12207 | -33.8677 | 6.582
0.14648 | -34.0407 | 4.1314
0.1709 | -34.0751 | 3.7251
0.19531 | -33.9816 | 3.1046
0.21973 | -33.9739 | 2.1177
0.24414 | -33.9283 | 1.8786
0.26855 | -33.9738 | 1.3021
0.29297 | -33.9853 | 1.2253
0.31738 | -33.9449 | 0.77877
0.3418 | -33.9447 | 0.97768
0.36621 | -33.9528 | 0.52354
0.39062 | -34.0448 | 0.23004
0.41504 | -33.9789 | 0.55136
0.43945 | -33.9654 | 0.42549
0.46387 | -34.0136 | -0.3117
0.48828 | -33.878 | -0.61272
0.5127 | -34.0318 | -0.071916
0.53711 | -34.0747 | 0.32333
0.56152 | -34.0289 | -0.38013
0.58594 | -33.9817 | 0.27439
0.61035 | -33.989 | -0.73839
0.63477 | -34.0139 | -0.45689
0.65918 | -33.9584 | -1.1828
0.68359 | -34.1128 | -0.94607
0.70801 | -34.0094 | -0.46538
0.73242 | -34.0396 | -0.98391
0.75684 | -34.0445 | -0.88407
0.78125 | -34.0259 | -1.0724
0.80566 | -34.1253 | -1.4499
0.83008 | -34.0574 | -0.79307
0.85449 | -33.9275 | -0.82816
0.87891 | -34.0904 | -1.2605
0.90332 | -33.9526 | -1.4846
0.92773 | -34.0384 | -1.0802
0.95215 | -34.0431 | -1.2598
0.97656 | -34.0302 | -0.71442
1.001 | -34.0026 | -1.1846
1.0254 | -33.9543 | -1.3895
1.0498 | -33.9652 | -1.8092
1.0742 | -34.1206 | -1.9762
1.0986 | -34.0411 | -1.787
1.123 | -34.0508 | -1.9603
1.1475 | -33.9907 | -1.6368
1.1719 | -34.0248 | -1.6413
1.1963 | -34.0322 | -2.0514
1.2207 | -34.0221 | -2.1663
1.2451 | -34.1298 | -1.7939
1.2695 | -34.0153 | -2.2661
1.2939 | -34.0238 | -2.4017
1.3184 | -33.9718 | -2.5763
1.3428 | -34.0814 | -1.9458
1.3672 | -33.9923 | -2.3309
1.3916 | -34.0756 | -2.7568
1.416 | -34.027 | -2.6115
1.4404 | -34.0548 | -2.8566
1.4648 | -33.9888 | -2.3715
1.4893 | -34.0204 | -2.6523
1.5137 | -34.068 | -2.1449
1.5381 | -34.0435 | -3.0058
1.5625 | -34.05 | -2.8755
1.5869 | -34.0525 | -3.4128
1.6113 | -33.9645 | -2.9043
1.6357 | -33.9639 | -3.0682
1.6602 | -34 | -3.4971
1.6846 | -34.0097 | -3.2851
1.709 | -34.0374 | -2.9786
1.7334 | -33.9816 | -3.2054
1.7578 | -34.021 | -4.1166
1.7822 | -34.0055 | -3.5944
1.8066 | -34.0564 | -4.2518
1.8311 | -34.0315 | -3.0545
1.8555 | -34.0527 | -4.1435
1.8799 | -34.1148 | -4.1135
1.9043 | -34.0664 | -4.1711
1.9287 | -33.9716 | -3.4237
1.9531 | -33.9868 | -4.0677
1.9775 | -34.019 | -3.6352
2.002 | -34.0107 | -3.8574
2.0264 | -34.0413 | -3.7492
2.0508 | -34.0417 | -4.4189
2.0752 | -34.0944 | -3.5523
2.0996 | -34.0774 | -3.965
2.124 | -34.1059 | -3.3504
2.1484 | -33.9894 | -3.733
2.1729 | -34.0234 | -4.376
2.1973 | -34.0442 | -4.1833
2.2217 | -34.0679 | -4.3603
2.2461 | -33.9336 | -3.5769
2.2705 | -34.1074 | -4.508
2.2949 | -34.0823 | -4.0512
2.3193 | -34.0128 | -4.4365
2.3438 | -34.012 | -4.3966
2.3682 | -34.0375 | -5.0055
2.3926 | -34.0294 | -4.6005
2.417 | -34.0129 | -4.6392
2.4414 | -34.0127 | -4.4241
2.4658 | -34.1249 | -4.8894
2.4902 | -34.0576 | -4.5601
2.5146 | -34.1363 | -5.0566
2.5391 | -34.0199 | -4.9699
2.5635 | -34.0195 | -5.166
2.5879 | -34.1447 | -4.7795
2.6123 | -33.945 | -4.7501
2.6367 | -34.1145 | -5.1081
2.6611 | -34.0718 | -5.2943
2.6855 | -34.0877 | -5.2678
2.71 | -33.9645 | -5.0019
2.7344 | -34.113 | -5.686
2.7588 | -34.0185 | -5.6132
2.7832 | -34.0714 | -5.61
2.8076 | -34.009 | -5.3428
2.832 | -34.0856 | -5.2184
2.8564 | -34.104 | -5.4739
2.8809 | -34.1017 | -5.863
2.9053 | -34.1051 | -5.5582
2.9297 | -34.1321 | -5.7148
2.9541 | -34.0653 | -5.7122
2.9785 | -34.0319 | -5.2702
3.0029 | -34.1133 | -5.7476
3.0273 | -34.1262 | -5.647
3.0518 | -34.0646 | -5.3027
3.0762 | -34.066 | -5.1634
3.1006 | -34.0135 | -5.9627
3.125 | -34.0609 | -6.0314
3.1494 | -34.1391 | -5.4854
3.1738 | -34.0858 | -5.9687
3.1982 | -34.1356 | -6.171
3.2227 | -34.0371 | -5.8627
3.2471 | -34.1364 | -6.1308
3.2715 | -34.1057 | -6.0724
3.2959 | -34.1727 | -5.9454
3.3203 | -34.1905 | -6.6005
3.3447 | -34.1228 | -6.1846
3.3691 | -34.08 | -5.6311
3.3936 | -34.1398 | -6.848
3.418 | -34.095 | -6.6047
3.4424 | -34.1754 | -6.3997
3.4668 | -34.1613 | -6.5785
3.4912 | -34.087 | -7.3114
3.5156 | -34.0529 | -7.5223
3.54 | -34.0889 | -7.589
3.5645 | -34.2118 | -6.4265
3.5889 | -34.0479 | -6.5018
3.6133 | -34.1332 | -7.5994
3.6377 | -34.1429 | -7.4374
3.6621 | -34.1055 | -7.057
3.6865 | -34.0142 | -6.7366
3.7109 | -34.1608 | -7.0533
3.7354 | -34.1399 | -6.6955
3.7598 | -34.0551 | -7.2379
3.7842 | -34.0284 | -7.0878
3.8086 | -34.0861 | -7.6389
3.833 | -34.0596 | -7.0485
3.8574 | -34.1566 | -7.2865
3.8818 | -34.153 | -7.061
3.9062 | -34.1177 | -7.4328
3.9307 | -34.034 | -7.0849
3.9551 | -34.056 | -7.7783
3.9795 | -34.098 | -7.9689
4.0039 | -34.1929 | -7.7336
4.0283 | -34.1886 | -7.3228
4.0527 | -34.1316 | -7.2475
4.0771 | -34.0937 | -7.666
4.1016 | -34.126 | -8.1804
4.126 | -34.0701 | -7.8105
4.1504 | -34.1272 | -8.405
4.1748 | -34.187 | -8.1103
4.1992 | -34.1228 | -7.9772
4.2236 | -34.0379 | -7.9968
4.248 | -34.055 | -8.3642
4.2725 | -34.1645 | -7.6918
4.2969 | -34.0762 | -7.546
4.3213 | -34.0361 | -7.9948
4.3457 | -34.0598 | -8.4182
4.3701 | -34.0492 | -8.9252
4.3945 | -34.1521 | -8.353
4.4189 | -34.0735 | -7.8632
4.4434 | -34.1279 | -8.2646
4.4678 | -34.1811 | -8.4923
4.4922 | -34.139 | -8.7059
4.5166 | -34.0754 | -8.4965
4.541 | -34.0922 | -8.5305
4.5654 | -34.0785 | -8.8307
4.5898 | -34.129 | -9.4869
4.6143 | -34.1868 | -8.7251
4.6387 | -34.1027 | -8.5259
4.6631 | -34.1105 | -8.7332
4.6875 | -34.1699 | -8.7372
4.7119 | -34.11 | -8.608
4.7363 | -34.1548 | -8.5847
4.7607 | -34.1143 | -8.7873
4.7852 | -34.1658 | -9.3833
4.8096 | -34.1346 | -9.0097
4.834 | -34.1512 | -9.2207
4.8584 | -34.1243 | -9.1584
4.8828 | -34.2212 | -9.3386
4.9072 | -34.1472 | -9.4694
4.9316 | -34.1804 | -9.7198
4.9561 | -34.2182 | -9.6321
4.9805 | -34.0684 | -9.1403
5.0049 | -34.1094 | -9.448
5.0293 | -34.1436 | -9.6512
5.0537 | -34.1187 | -10.0087
5.0781 | -34.1997 | -9.8248
5.1025 | -34.1372 | -9.8255
5.127 | -34.146 | -9.8251
5.1514 | -34.2035 | -10.1687
5.1758 | -34.1951 | -9.8211
5.2002 | -34.173 | -9.6594
5.2246 | -34.2052 | -9.6919
5.249 | -34.0539 | -9.9801
5.2734 | -34.1075 | -10.1617
5.2979 | -34.1864 | -9.9686
5.3223 | -34.1804 | -10.3851
5.3467 | -34.1841 | -9.8828
5.3711 | -34.1694 | -10.0602
5.3955 | -34.0993 | -10.1721
5.4199 | -34.2044 | -10.6676
5.4443 | -34.1201 | -9.9598
5.4688 | -34.1406 | -10.0198
5.4932 | -34.157 | -10.5164
5.5176 | -34.1882 | -10.2461
5.542 | -34.2384 | -10.4247
5.5664 | -34.2154 | -11.013
5.5908 | -34.0668 | -9.7518
5.6152 | -34.1626 | -11.1594
5.6396 | -34.1576 | -10.7134
5.6641 | -34.1676 | -10.3541
5.6885 | -34.1713 | -10.1928
5.7129 | -34.1843 | -11.0545
5.7373 | -34.1942 | -10.4301
5.7617 | -34.1438 | -10.7668
5.7861 | -34.2086 | -10.9483
5.8105 | -34.2136 | -11.0548
5.835 | -34.1643 | -10.4994
5.8594 | -34.2144 | -11.1213
5.8838 | -34.2508 | -11.0402
5.9082 | -34.1774 | -11.354
5.9326 | -34.1668 | -11.1601
5.957 | -34.1649 | -11.5629
5.9814 | -34.3003 | -11.1445
6.0059 | -34.1542 | -11.2243
6.0303 | -34.1553 | -11.7354
6.0547 | -34.1934 | -11.3396
6.0791 | -34.1711 | -11.5934
6.1035 | -34.1476 | -11.2642
6.1279 | -34.1364 | -11.308
6.1523 | -34.1585 | -11.6801
6.1768 | -34.1545 | -11.4251
6.2012 | -34.169 | -11.7021
6.2256 | -34.2146 | -11.7529
6.25 | -34.197 | -11.7668
6.2744 | -34.1478 | -12.0838
6.2988 | -34.2899 | -12.0531
6.3232 | -34.1695 | -11.9961
6.3477 | -34.2862 | -11.5522
6.3721 | -34.2096 | -11.9898
6.3965 | -34.2694 | -12.0724
6.4209 | -34.1527 | -12.1005
6.4453 | -34.1355 | -12.05
6.4697 | -34.2147 | -12.2268
6.4941 | -34.2605 | -12.3012
6.5186 | -34.2882 | -12.7049
6.543 | -34.2666 | -12.4275
6.5674 | -34.1538 | -12.398
6.5918 | -34.2273 | -13.1645
6.6162 | -34.1968 | -12.9498
6.6406 | -34.1807 | -12.6643
6.665 | -34.3244 | -12.6068
6.6895 | -34.2664 | -12.5434
6.7139 | -34.3417 | -12.9235
6.7383 | -34.2142 | -12.6436
6.7627 | -34.226 | -12.0024
6.7871 | -34.2335 | -12.8266
6.8115 | -34.2035 | -12.5568
6.8359 | -34.268 | -12.7161
6.8604 | -34.2673 | -12.8581
6.8848 | -34.1997 | -13.2756
6.9092 | -34.2953 | -12.824
6.9336 | -34.235 | -12.2754
6.958 | -34.2732 | -13.6316
6.9824 | -34.2204 | -13.289
7.0068 | -34.2361 | -13.1437
7.0312 | -34.2484 | -12.9517
7.0557 | -34.1608 | -13.0023
7.0801 | -34.2227 | -13.0298
7.1045 | -34.2364 | -13.9482
7.1289 | -34.2425 | -13.4984
7.1533 | -34.1851 | -13.2334
7.1777 | -34.2527 | -13.7286
7.2021 | -34.2746 | -13.7684
7.2266 | -34.2732 | -13.2637
7.251 | -34.1963 | -13.693
7.2754 | -34.249 | -13.4911
7.2998 | -34.225 | -13.9573
7.3242 | -34.2804 | -13.5578
7.3486 | -34.2452 | -13.9016
7.373 | -34.3023 | -14.0014
7.3975 | -34.3098 | -13.438
7.4219 | -34.3693 | -14.1285
7.4463 | -34.244 | -13.6621
7.4707 | -34.295 | -13.9496
7.4951 | -34.2714 | -13.7228
7.5195 | -34.2306 | -14.1854
7.5439 | -34.3141 | -13.8155
7.5684 | -34.2113 | -13.9008
7.5928 | -34.3224 | -14.6156
7.6172 | -34.3174 | -13.9337
7.6416 | -34.2987 | -13.8858
7.666 | -34.2793 | -14.3241
7.6904 | -34.4614 | -14.6844
7.7148 | -34.3419 | -14.3703
7.7393 | -34.3833 | -14.3243
7.7637 | -34.3146 | -14.5369
7.7881 | -34.3275 | -15.0053
7.8125 | -34.2856 | -14.9006
7.8369 | -34.3182 | -14.6483
7.8613 | -34.3587 | -14.4227
7.8857 | -34.2508 | -14.7759
7.9102 | -34.2577 | -14.4831
7.9346 | -34.3458 | -15.0786
7.959 | -34.3958 | -14.6392
7.9834 | -34.2695 | -14.6715
8.0078 | -34.2853 | -14.9679
8.0322 | -34.2864 | -15.1711
8.0566 | -34.3402 | -15.3412
8.0811 | -34.2979 | -14.9805
8.1055 | -34.349 | -15.196
8.1299 | -34.2868 | -15.0984
8.1543 | -34.2685 | -15.1192
8.1787 | -34.3018 | -14.9989
8.2031 | -34.3331 | -14.8579
8.2275 | -34.2543 | -14.7375
8.252 | -34.2827 | -15.5176
8.2764 | -34.3306 | -15.2699
8.3008 | -34.2929 | -15.6719
8.3252 | -34.3246 | -15.6339
8.3496 | -34.2616 | -15.2669
8.374 | -34.3525 | -15.615
8.3984 | -34.3459 | -15.1523
8.4229 | -34.2426 | -16.1395
8.4473 | -34.343 | -16.1774
8.4717 | -34.437 | -15.8726
8.4961 | -34.3406 | -16.005
8.5205 | -34.3705 | -15.5046
8.5449 | -34.267 | -15.6649
8.5693 | -34.2949 | -16.1954
8.5938 | -34.3329 | -16.1009
8.6182 | -34.2998 | -16.0106
8.6426 | -34.3885 | -15.7638
8.667 | -34.2937 | -15.7826
8.6914 | -34.3623 | -16.1162
8.7158 | -34.302 | -15.9148
8.7402 | -34.3421 | -16.2174
8.7646 | -34.2881 | -15.6837
8.7891 | -34.3463 | -16.0709
8.8135 | -34.3267 | -16.7348
8.8379 | -34.2403 | -16.0933
8.8623 | -34.3276 | -16.5102
8.8867 | -34.4624 | -16.3998
8.9111 | -34.4139 | -16.6629
8.9355 | -34.3859 | -16.4072
8.96 | -34.3607 | -17.1952
8.9844 | -34.3792 | -16.7491
9.0088 | -34.3726 | -16.6951
9.0332 | -34.473 | -16.3312
9.0576 | -34.3448 | -16.5825
9.082 | -34.3453 | -16.5738
9.1064 | -34.4042 | -16.5108
9.1309 | -34.3698 | -16.9613
9.1553 | -34.4001 | -17.2311
9.1797 | -34.3567 | -16.6775
9.2041 | -34.4492 | -16.5733
9.2285 | -34.3663 | -16.9146
9.2529 | -34.3488 | -16.695
9.2773 | -34.3465 | -17.1822
9.3018 | -34.3684 | -17.1515
9.3262 | -34.4433 | -17.0206
9.3506 | -34.4659 | -17.4875
9.375 | -34.389 | -17.0218
9.3994 | -34.3613 | -17.4909
9.4238 | -34.3479 | -17.6082
9.4482 | -34.3771 | -17.2119
9.4727 | -34.3393 | -17.7913
9.4971 | -34.3297 | -18.1972
9.5215 | -34.4468 | -17.9051
9.5459 | -34.4359 | -17.4952
9.5703 | -34.5046 | -17.4045
9.5947 | -34.3614 | -17.9336
9.6191 | -34.4558 | -16.8935
9.6436 | -34.3414 | -17.7647
9.668 | -34.381 | -18.1717
9.6924 | -34.4334 | -17.9209
9.7168 | -34.4432 | -17.9576
9.7412 | -34.3761 | -18.4518
9.7656 | -34.4754 | -17.7965
9.79 | -34.3595 | -18.2708
9.8145 | -34.4512 | -17.6622
9.8389 | -34.4161 | -17.775
9.8633 | -34.4186 | -17.7388
9.8877 | -34.4623 | -18.2707
9.9121 | -34.3944 | -18.3418
9.9365 | -34.3534 | -18.2056
9.9609 | -34.4132 | -18.1341
9.9854 | -34.4524 | -18.0279
10.0098 | -34.4105 | -18.4668

## ACOUS_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.1476 | 9.4956
0.097656 | -34.1699 | 4.723
0.12207 | -33.9643 | 3.6681
0.14648 | -34.0757 | 1.3972
0.1709 | -34.0019 | 1.2564
0.19531 | -34.0262 | 0.24926
0.21973 | -34.0114 | 0.35143
0.24414 | -34.0643 | -0.91166
0.26855 | -34.0685 | -0.12269
0.29297 | -34.0177 | -1.4597
0.31738 | -34.0512 | -0.99144
0.3418 | -34.0118 | -1.6437
0.36621 | -34.0828 | -1.3249
0.39062 | -33.9809 | -2.0641
0.41504 | -34.1044 | -1.6789
0.43945 | -34.0083 | -2.1018
0.46387 | -34.1025 | -2.2684
0.48828 | -33.9766 | -2.2424
0.5127 | -34.1201 | -2.5775
0.53711 | -34.0101 | -2.2956
0.56152 | -34.1015 | -2.9401
0.58594 | -34.0145 | -2.4047
0.61035 | -34.0873 | -3.1963
0.63477 | -34.0369 | -2.5264
0.65918 | -34.0788 | -3.4144
0.68359 | -34.0391 | -2.6223
0.70801 | -34.0358 | -3.4783
0.73242 | -34.0792 | -2.858
0.75684 | -34.0493 | -3.6513
0.78125 | -34.0967 | -2.9961
0.80566 | -34.0182 | -3.758
0.83008 | -34.1182 | -3.3399
0.85449 | -33.9948 | -3.7863
0.87891 | -34.0823 | -3.5096
0.90332 | -33.9835 | -3.5584
0.92773 | -34.128 | -3.5632
0.95215 | -34.0248 | -3.5304
0.97656 | -34.1334 | -4.0397
1.001 | -34.024 | -3.6532
1.0254 | -34.1271 | -4.222
1.0498 | -34.0685 | -3.7021
1.0742 | -34.1231 | -4.5794
1.0986 | -34.0478 | -3.8406
1.123 | -34.0692 | -4.5758
1.1475 | -34.0777 | -3.8255
1.1719 | -34.0733 | -4.7151
1.1963 | -34.1039 | -3.8268
1.2207 | -34.0568 | -4.8297
1.2451 | -34.1397 | -4.0636
1.2695 | -34.038 | -4.7219
1.2939 | -34.1375 | -4.237
1.3184 | -34.0105 | -4.6232
1.3428 | -34.1682 | -4.4115
1.3672 | -34.0421 | -4.5996
1.3916 | -34.1464 | -4.765
1.416 | -34.0094 | -4.683
1.4404 | -34.1371 | -4.8786
1.4648 | -34.0202 | -4.3663
1.4893 | -34.1383 | -4.9157
1.5137 | -34.067 | -4.4168
1.5381 | -34.1356 | -5.33
1.5625 | -34.0962 | -4.5115
1.5869 | -34.1212 | -5.3755
1.6113 | -34.1217 | -4.7098
1.6357 | -34.0665 | -5.6486
1.6602 | -34.1125 | -4.8875
1.6846 | -34.0549 | -5.4645
1.709 | -34.1397 | -4.9211
1.7334 | -34.0243 | -5.5025
1.7578 | -34.1685 | -5.2127
1.7822 | -34.0248 | -5.3406
1.8066 | -34.1594 | -5.273
1.8311 | -34.0329 | -5.2818
1.8555 | -34.1902 | -5.4639
1.8799 | -34.0578 | -5.3074
1.9043 | -34.157 | -5.9147
1.9287 | -34.0383 | -5.2961
1.9531 | -34.1346 | -5.9929
1.9775 | -34.0749 | -5.114
2.002 | -34.125 | -6.0524
2.0264 | -34.0939 | -5.2648
2.0508 | -34.1033 | -6.0007
2.0752 | -34.1375 | -5.3367
2.0996 | -34.1133 | -6.1738
2.124 | -34.1587 | -5.4783
2.1484 | -34.09 | -6.3861
2.1729 | -34.1629 | -5.8156
2.1973 | -34.0677 | -6.1449
2.2217 | -34.1687 | -5.9442
2.2461 | -34.0687 | -6.1352
2.2705 | -34.2178 | -6.1463
2.2949 | -34.0794 | -6.1835
2.3193 | -34.1812 | -6.3685
2.3438 | -34.081 | -6.0212
2.3682 | -34.1498 | -6.5391
2.3926 | -34.0749 | -5.8796
2.417 | -34.1415 | -6.6209
2.4414 | -34.1107 | -5.8155
2.4658 | -34.1726 | -6.6734
2.4902 | -34.1595 | -5.9451
2.5146 | -34.1347 | -6.9224
2.5391 | -34.1476 | -6.1851
2.5635 | -34.1146 | -6.8751
2.5879 | -34.17 | -6.2889
2.6123 | -34.0896 | -6.7708
2.6367 | -34.1794 | -6.4267
2.6611 | -34.0872 | -6.8618
2.6855 | -34.2099 | -6.5688
2.71 | -34.0653 | -6.796
2.7344 | -34.2171 | -6.8178
2.7588 | -34.0992 | -6.4868
2.7832 | -34.1887 | -7.1165
2.8076 | -34.0806 | -6.5669
2.832 | -34.1902 | -7.1064
2.8564 | -34.1353 | -6.7027
2.8809 | -34.1625 | -7.5631
2.9053 | -34.1173 | -6.6847
2.9297 | -34.1607 | -7.5036
2.9541 | -34.1667 | -6.7997
2.9785 | -34.1199 | -7.5394
3.0029 | -34.1846 | -6.8188
3.0273 | -34.1233 | -7.5323
3.0518 | -34.2129 | -6.8839
3.0762 | -34.1054 | -7.4592
3.1006 | -34.2145 | -7.0866
3.125 | -34.102 | -7.4287
3.1494 | -34.2167 | -7.3251
3.1738 | -34.1071 | -7.5086
3.1982 | -34.2209 | -7.6167
3.2227 | -34.0829 | -7.3644
3.2471 | -34.2045 | -7.7737
3.2715 | -34.122 | -7.2791
3.2959 | -34.2079 | -7.7223
3.3203 | -34.1371 | -7.1545
3.3447 | -34.1714 | -7.9559
3.3691 | -34.1616 | -7.288
3.3936 | -34.1778 | -8.0252
3.418 | -34.1818 | -7.3896
3.4424 | -34.1521 | -8.2258
3.4668 | -34.2248 | -7.4864
3.4912 | -34.1432 | -8.3324
3.5156 | -34.2581 | -7.7288
3.54 | -34.1044 | -8.1894
3.5645 | -34.2256 | -7.9132
3.5889 | -34.1389 | -8.1386
3.6133 | -34.2447 | -8.0185
3.6377 | -34.1129 | -8.1436
3.6621 | -34.2324 | -8.2778
3.6865 | -34.1283 | -7.8586
3.7109 | -34.2411 | -8.3362
3.7354 | -34.1694 | -7.9875
3.7598 | -34.2235 | -8.564
3.7842 | -34.1732 | -8.1001
3.8086 | -34.1871 | -8.657
3.833 | -34.1756 | -8.0121
3.8574 | -34.1807 | -8.6952
3.8818 | -34.235 | -8.0661
3.9062 | -34.1676 | -8.8509
3.9307 | -34.2234 | -8.0558
3.9551 | -34.1611 | -8.7078
3.9795 | -34.279 | -8.3981
4.0039 | -34.1505 | -8.7119
4.0283 | -34.26 | -8.5309
4.0527 | -34.1163 | -8.7643
4.0771 | -34.2715 | -8.8171
4.1016 | -34.1705 | -8.6246
4.126 | -34.2693 | -8.9222
4.1504 | -34.1596 | -8.6751
4.1748 | -34.261 | -9.2425
4.1992 | -34.1809 | -8.6687
4.2236 | -34.2145 | -9.3539
4.248 | -34.1812 | -8.6129
4.2725 | -34.1978 | -9.2991
4.2969 | -34.2008 | -8.4718
4.3213 | -34.216 | -9.4041
4.3457 | -34.2231 | -8.6904
4.3701 | -34.1718 | -9.4124
4.3945 | -34.247 | -8.8067
4.4189 | -34.1723 | -9.3641
4.4434 | -34.2891 | -8.9268
4.4678 | -34.1939 | -9.4072
4.4922 | -34.2812 | -9.4082
4.5166 | -34.1405 | -9.3506
4.541 | -34.2816 | -9.4634
4.5654 | -34.1651 | -9.2316
4.5898 | -34.2811 | -9.7311
4.6143 | -34.169 | -9.1124
4.6387 | -34.2553 | -9.8747
4.6631 | -34.205 | -9.2032
4.6875 | -34.2486 | -9.9705
4.7119 | -34.2107 | -9.2484
4.7363 | -34.2299 | -9.7526
4.7607 | -34.2402 | -9.2557
4.7852 | -34.2341 | -10.0701
4.8096 | -34.2655 | -9.432
4.834 | -34.1961 | -10.1289
4.8584 | -34.3058 | -9.3636
4.8828 | -34.1808 | -9.8232
4.9072 | -34.3111 | -9.6215
4.9316 | -34.1815 | -10.0368
4.9561 | -34.3078 | -9.9807
4.9805 | -34.1762 | -9.8205
5.0049 | -34.2853 | -10.1915
5.0293 | -34.1471 | -9.7927
5.0537 | -34.269 | -10.342
5.0781 | -34.1851 | -9.6289
5.1025 | -34.2822 | -10.4638
5.127 | -34.2153 | -9.7713
5.1514 | -34.2783 | -10.6211
5.1758 | -34.2518 | -9.8008
5.2002 | -34.2607 | -10.6262
5.2246 | -34.2743 | -9.8199
5.249 | -34.2085 | -10.5992
5.2734 | -34.2859 | -9.8294
5.2979 | -34.2243 | -10.7427
5.3223 | -34.3377 | -10.1997
5.3467 | -34.1954 | -10.5127
5.3711 | -34.3282 | -10.3674
5.3955 | -34.1812 | -10.5858
5.4199 | -34.3466 | -10.6583
5.4443 | -34.2077 | -10.4538
5.4688 | -34.3061 | -10.9421
5.4932 | -34.1881 | -10.3616
5.5176 | -34.3034 | -10.8862
5.542 | -34.2268 | -10.4552
5.5664 | -34.2931 | -11.1153
5.5908 | -34.2445 | -10.4787
5.6152 | -34.2731 | -11.2306
5.6396 | -34.256 | -10.6826
5.6641 | -34.2411 | -11.3185
5.6885 | -34.289 | -10.5652
5.7129 | -34.2153 | -11.3256
5.7373 | -34.2923 | -10.6835
5.7617 | -34.242 | -11.182
5.7861 | -34.3646 | -10.573
5.8105 | -34.2121 | -11.1913
5.835 | -34.3772 | -11.0079
5.8594 | -34.2113 | -11.1057
5.8838 | -34.3498 | -11.1716
5.9082 | -34.2256 | -11.1259
5.9326 | -34.3423 | -11.5117
5.957 | -34.2437 | -10.8493
5.9814 | -34.3366 | -11.6033
6.0059 | -34.2583 | -11.0412
6.0303 | -34.3047 | -11.7451
6.0547 | -34.2611 | -11.0252
6.0791 | -34.2741 | -11.5968
6.1035 | -34.2931 | -10.7926
6.1279 | -34.2693 | -11.8774
6.1523 | -34.3185 | -11.1774
6.1768 | -34.2399 | -11.9598
6.2012 | -34.3294 | -11.4284
6.2256 | -34.2362 | -11.6703
6.25 | -34.3698 | -11.3515
6.2744 | -34.2532 | -11.5795
6.2988 | -34.3446 | -11.7272
6.3232 | -34.2482 | -11.5082
6.3477 | -34.3634 | -11.6112
6.3721 | -34.2332 | -11.3891
6.3965 | -34.3525 | -11.9717
6.4209 | -34.2546 | -11.513
6.4453 | -34.3354 | -12.0877
6.4697 | -34.2862 | -11.4236
6.4941 | -34.328 | -12.4079
6.5186 | -34.3284 | -11.5252
6.543 | -34.3254 | -12.5026
6.5674 | -34.3555 | -11.5847
6.5918 | -34.3001 | -12.5472
6.6162 | -34.3409 | -11.866
6.6406 | -34.2416 | -12.3877
6.665 | -34.3906 | -11.9243
6.6895 | -34.2564 | -12.1504
6.7139 | -34.3856 | -12.1976
6.7383 | -34.2488 | -12.265
6.7627 | -34.3604 | -12.3945
6.7871 | -34.2297 | -12.1389
6.8115 | -34.4075 | -12.4244
6.8359 | -34.2975 | -12.3504
6.8604 | -34.3801 | -12.6965
6.8848 | -34.2706 | -12.0933
6.9092 | -34.3625 | -12.7795
6.9336 | -34.3443 | -12.2285
6.958 | -34.3385 | -12.9516
6.9824 | -34.3493 | -12.135
7.0068 | -34.333 | -12.9061
7.0312 | -34.3581 | -12.1293
7.0557 | -34.3194 | -12.8477
7.0801 | -34.3849 | -12.2119
7.1045 | -34.3248 | -12.7444
7.1289 | -34.4253 | -12.4852
7.1533 | -34.2939 | -12.8807
7.1777 | -34.4162 | -12.6601
7.2021 | -34.3119 | -12.8912
7.2266 | -34.4446 | -12.8311
7.251 | -34.3432 | -12.8869
7.2754 | -34.4201 | -13.0247
7.2998 | -34.2827 | -12.4661
7.3242 | -34.3817 | -13.2779
7.3486 | -34.3375 | -12.5158
7.373 | -34.3815 | -13.3291
7.3975 | -34.3593 | -12.5662
7.4219 | -34.3833 | -13.5296
7.4463 | -34.3684 | -12.6858
7.4707 | -34.3359 | -13.4226
7.4951 | -34.3603 | -12.8777
7.5195 | -34.3313 | -13.3666
7.5439 | -34.4031 | -12.9603
7.5684 | -34.3074 | -13.2756
7.5928 | -34.3916 | -13.071
7.6172 | -34.283 | -13.2557
7.6416 | -34.4483 | -13.2789
7.666 | -34.3018 | -13.0099
7.6904 | -34.4655 | -13.5425
7.7148 | -34.3479 | -13.2053
7.7393 | -34.4369 | -13.6762
7.7637 | -34.3906 | -13.2498
7.7881 | -34.4349 | -14.0163
7.8125 | -34.3745 | -13.3154
7.8369 | -34.377 | -14.0033
7.8613 | -34.3777 | -13.1074
7.8857 | -34.3738 | -14.21
7.9102 | -34.4326 | -13.2516
7.9346 | -34.3354 | -14.094
7.959 | -34.4447 | -13.42
7.9834 | -34.3513 | -14.0877
8.0078 | -34.4638 | -13.5526
8.0322 | -34.3605 | -14.0278
8.0566 | -34.4735 | -13.8077
8.0811 | -34.3535 | -13.7299
8.1055 | -34.4648 | -13.7779
8.1299 | -34.3255 | -13.7123
8.1543 | -34.4451 | -14.2019
8.1787 | -34.3563 | -13.6659
8.2031 | -34.4115 | -14.0977
8.2275 | -34.3866 | -13.6803
8.252 | -34.4264 | -14.404
8.2764 | -34.4148 | -13.8569
8.3008 | -34.4189 | -14.4119
8.3252 | -34.4278 | -13.8408
8.3496 | -34.4123 | -14.6325
8.374 | -34.4675 | -13.7563
8.3984 | -34.3863 | -14.8004
8.4229 | -34.4433 | -13.9633
8.4473 | -34.3947 | -14.5783
8.4717 | -34.46 | -14.0703
8.4961 | -34.3581 | -14.6296
8.5205 | -34.4895 | -14.0459
8.5449 | -34.3539 | -14.5286
8.5693 | -34.519 | -14.4111
8.5938 | -34.3669 | -14.1109
8.6182 | -34.5009 | -14.839
8.6426 | -34.3868 | -14.1478
8.667 | -34.4583 | -14.9221
8.6914 | -34.3926 | -13.9048
8.7158 | -34.4297 | -15.0984
8.7402 | -34.4019 | -14.1597
8.7646 | -34.43 | -14.9464
8.7891 | -34.4728 | -14.3739
8.8135 | -34.4219 | -15.2679
8.8379 | -34.5046 | -14.2187
8.8623 | -34.4197 | -14.8766
8.8867 | -34.5239 | -14.697
8.9111 | -34.3966 | -14.792
8.9355 | -34.5032 | -14.8924
8.96 | -34.3993 | -15.0688
8.9844 | -34.5159 | -15.0694
9.0088 | -34.4057 | -14.9919
9.0332 | -34.5208 | -15.0514
9.0576 | -34.3661 | -14.829
9.082 | -34.5402 | -15.2863
9.1064 | -34.4336 | -14.5454
9.1309 | -34.5075 | -15.4091
9.1553 | -34.4267 | -14.7356
9.1797 | -34.4653 | -15.5971
9.2041 | -34.4636 | -14.8664
9.2285 | -34.4625 | -15.5597
9.2529 | -34.4476 | -14.6914
9.2773 | -34.4184 | -15.8094
9.3018 | -34.4979 | -14.8321
9.3262 | -34.4532 | -15.501
9.3506 | -34.558 | -15.1391
9.375 | -34.4108 | -15.3582
9.3994 | -34.5291 | -15.0428
9.4238 | -34.4385 | -15.3272
9.4482 | -34.5751 | -15.1564
9.4727 | -34.429 | -15.3759
9.4971 | -34.5316 | -15.6983
9.5215 | -34.4136 | -15.267
9.5459 | -34.5058 | -15.968
9.5703 | -34.455 | -15.1096
9.5947 | -34.4951 | -16.0721
9.6191 | -34.4625 | -15.304
9.6436 | -34.5209 | -15.9913
9.668 | -34.5005 | -15.2785
9.6924 | -34.44 | -16.0106
9.7168 | -34.5117 | -15.3698
9.7412 | -34.4599 | -16.1026
9.7656 | -34.5512 | -15.3141
9.79 | -34.4211 | -16.0986
9.8145 | -34.5609 | -15.4314
9.8389 | -34.4323 | -16.3561
9.8633 | -34.5665 | -15.7688
9.8877 | -34.4176 | -15.872
9.9121 | -34.5543 | -15.8564
9.9365 | -34.4385 | -16.0064
9.9609 | -34.5789 | -16.0171
9.9854 | -34.4706 | -15.8429
10.0098 | -34.5931 | -16.5139

## SEALED_NOISE
dBPa/√Hz | Frequency
0.009769 | -31.8736
0.01013 | -29.3588
0.010492 | -29.4025
0.010854 | -29.9998
0.011216 | -30.5629
0.011578 | -32.1224
0.011939 | -34.9612
0.012301 | -37.5183
0.012663 | -37.2943
0.013025 | -36.27
0.013387 | -37.7109
0.013748 | -38.3953
0.01411 | -38.6457
0.014472 | -38.3606
0.014834 | -37.8058
0.015196 | -39.4364
0.015557 | -38.1606
0.015919 | -35.7226
0.016281 | -35.1792
0.016643 | -33.7622
0.017005 | -33.2687
0.017366 | -31.9893
0.017728 | -31.993
0.01809 | -34.4934
0.018452 | -36.7347
0.018814 | -36.4838
0.019175 | -36.8369
0.019537 | -36.2219
0.019899 | -34.5641
0.020261 | -33.5524
0.020623 | -34.487
0.020984 | -36.089
0.021346 | -38.5003
0.021708 | -39.7387
0.02207 | -38.872
0.022432 | -37.8583
0.022793 | -37.2092
0.023155 | -36.977
0.023517 | -37.4068
0.023879 | -37.912
0.024241 | -37.4494
0.024602 | -36.8167
0.024964 | -36.7246
0.025326 | -36.7617
0.025688 | -36.7204
0.02605 | -36.9681
0.026411 | -37.7697
0.026773 | -38.9509
0.027135 | -39.5911
0.027497 | -39.4662
0.027859 | -39.8121
0.02822 | -41.3638
0.028582 | -43.5416
0.028944 | -44.5686
0.029306 | -43.7012
0.029668 | -41.9662
0.030029 | -40.5652
0.030391 | -40.224
0.030753 | -40.9568
0.031115 | -41.9287
0.031477 | -42.7202
0.031838 | -42.843
0.0322 | -41.9703
0.032562 | -40.5929
0.032924 | -39.5306
0.033286 | -39.2026
0.033647 | -39.7304
0.034009 | -41.0203
0.034371 | -42.7513
0.034733 | -44.5259
0.035095 | -46.1175
0.035457 | -47.1826
0.035818 | -47.4097
0.03618 | -46.9397
0.036542 | -46.2588
0.036904 | -45.845
0.037266 | -45.724
0.037627 | -45.7127
0.037989 | -45.9025
0.038351 | -46.4797
0.038713 | -47.1063
0.039075 | -47.1389
0.039436 | -46.5982
0.039798 | -46.0839
0.04016 | -45.5961
0.040522 | -45.0105
0.040884 | -44.3268
0.041245 | -43.8087
0.041607 | -43.9817
0.041969 | -44.166
0.042331 | -44.5716
0.042693 | -45.0608
0.043054 | -45.5535
0.043416 | -46.0148
0.043778 | -46.5495
0.04414 | -47.1337
0.044502 | -47.7291
0.044863 | -48.2874
0.045225 | -48.7037
0.045587 | -49.0876
0.045949 | -49.4962
0.046311 | -49.798
0.046672 | -50.0344
0.047034 | -50.1421
0.047396 | -50.0388
0.047758 | -49.8184
0.04812 | -49.5521
0.048481 | -49.3017
0.048843 | -48.9107
0.049205 | -48.4351
0.049567 | -48.0434
0.049929 | -47.9429
0.05029 | -48.2388
0.050652 | -48.732
0.051014 | -49.2713
0.051376 | -49.7924
0.051738 | -50.3517
0.052099 | -50.8715
0.052461 | -51.4303
0.052823 | -51.7885
0.053185 | -51.9943
0.053547 | -52.0215
0.053908 | -51.9382
0.05427 | -51.8099
0.054632 | -51.5766
0.054994 | -51.2126
0.055356 | -50.7502
0.055717 | -50.0904
0.056079 | -49.2819
0.056441 | -48.5218
0.056803 | -47.998
0.057165 | -47.8477
0.057526 | -48.1093
0.057888 | -48.7208
0.05825 | -49.56
0.058612 | -50.3088
0.058974 | -50.6818
0.059335 | -50.7247
0.059697 | -50.7215
0.060059 | -50.856
0.060421 | -51.3347
0.060783 | -52.1758
0.061144 | -53.1341
0.061506 | -53.8184
0.061868 | -53.9187
0.06223 | -53.5631
0.062592 | -53.163
0.062953 | -53.0342
0.063315 | -52.9666
0.063677 | -53.0686
0.064039 | -53.3623
0.064401 | -53.8064
0.064762 | -54.3512
0.065124 | -55.0159
0.065486 | -55.7916
0.065848 | -56.5962
0.06621 | -57.3639
0.066571 | -58.0025
0.066933 | -58.5163
0.067295 | -58.5495
0.067657 | -58.256
0.068019 | -57.6382
0.06838 | -56.7917
0.068742 | -55.977
0.069104 | -55.2771
0.069466 | -54.7552
0.069828 | -54.5328
0.070189 | -54.579
0.070551 | -54.8082
0.070913 | -55.0777
0.071275 | -55.316
0.071637 | -55.4267
0.071998 | -55.3475
0.07236 | -55.1333
0.072722 | -54.7752
0.073084 | -54.5799
0.073446 | -54.3362
0.073807 | -54.0907
0.074169 | -53.9256
0.074531 | -53.8017
0.074893 | -53.7919
0.075255 | -53.9082
0.075616 | -54.0904
0.075978 | -54.3561
0.07634 | -54.673
0.076702 | -54.9849
0.077064 | -55.1734
0.077425 | -55.2617
0.077787 | -55.2763
0.078149 | -55.2204
0.078511 | -55.1979
0.078873 | -55.2019
0.079234 | -55.2233
0.079596 | -55.3475
0.079958 | -55.5472
0.08032 | -55.7083
0.080682 | -55.7353
0.081043 | -55.6567
0.081405 | -55.4799
0.081767 | -55.1836
0.082129 | -54.791
0.082491 | -54.3657
0.082852 | -54.0092
0.083214 | -54.0347
0.083576 | -54.0364
0.083938 | -54.181
0.0843 | -54.4486
0.084661 | -54.8658
0.085023 | -55.3448
0.085385 | -55.7982
0.085747 | -56.1342
0.086109 | -56.3278
0.08647 | -56.3723
0.086832 | -56.3047
0.087194 | -56.2105
0.087556 | -55.9843
0.087918 | -55.7704
0.088279 | -55.6576
0.088641 | -55.6821
0.089003 | -55.8443
0.089365 | -56.1276
0.089727 | -56.5303
0.090088 | -56.9764
0.09045 | -57.4213
0.090812 | -57.8367
0.091174 | -58.1167
0.091536 | -58.1412
0.091897 | -57.9894
0.092259 | -57.7441
0.092621 | -57.4498
0.092983 | -57.0448
0.093345 | -56.6796
0.093706 | -56.5687
0.094068 | -56.4527
0.09443 | -56.4429
0.094792 | -56.5611
0.095154 | -56.8035
0.095515 | -57.1863
0.095877 | -57.6729
0.096239 | -58.2663
0.096601 | -58.9413
0.096963 | -59.6103
0.097324 | -60.2735
0.097686 | -60.789
0.098048 | -61.1612
0.09841 | -61.4501
0.098772 | -61.5428
0.099133 | -61.4092
0.099495 | -61.1681
0.099857 | -60.8987
0.10022 | -60.5388
0.10058 | -60.1395
0.10094 | -59.6921
0.1013 | -59.234
0.10167 | -58.7966
0.10203 | -58.3967
0.10239 | -58.051
0.10275 | -57.7912
0.10311 | -57.6005
0.10347 | -57.4996
0.10384 | -57.5053
0.1042 | -57.795
0.10456 | -57.9828
0.10492 | -58.2651
0.10528 | -58.622
0.10565 | -59.0198
0.10601 | -59.3742
0.10637 | -59.7157
0.10673 | -59.9901
0.10709 | -60.1844
0.10745 | -60.2122
0.10782 | -59.9959
0.10818 | -59.6487
0.10854 | -59.2539
0.1089 | -58.8552
0.10926 | -58.4634
0.10963 | -58.0948
0.10999 | -57.7788
0.11035 | -57.5423
0.11071 | -57.3747
0.11107 | -57.2949
0.11144 | -57.3165
0.1118 | -57.4372
0.11216 | -57.6566
0.11252 | -57.9853
0.11288 | -58.4228
0.11324 | -58.9642
0.11361 | -59.6185
0.11397 | -60.3815
0.11433 | -61.2344
0.11469 | -61.7907
0.11505 | -62.5791
0.11542 | -63.3668
0.11578 | -64.0786
0.11614 | -64.6797
0.1165 | -65.2778
0.11686 | -65.8375
0.11722 | -66.0073
0.11758 | -65.8543
0.11795 | -65.6432
0.11831 | -65.3775
0.11867 | -65.0898
0.11903 | -64.786
0.11939 | -64.4943
0.11976 | -64.1996
0.12012 | -63.8856
0.12048 | -63.5642
0.12084 | -63.3028
0.1212 | -63.0893
0.12157 | -62.9499
0.12193 | -62.8585
0.12229 | -62.8416
0.12265 | -62.8214
0.12301 | -62.844
0.12337 | -62.9217
0.12374 | -62.9853
0.1241 | -62.9904
0.12446 | -62.9535
0.12482 | -62.8832
0.12518 | -62.6932
0.12554 | -62.517
0.12591 | -62.2764
0.12627 | -62.021
0.12663 | -61.7708
0.12699 | -61.5321
0.12735 | -61.237
0.12772 | -60.9526
0.12808 | -60.6919
0.12844 | -60.4384
0.1288 | -60.2072
0.12916 | -59.9919
0.12953 | -59.818
0.12989 | -59.7
0.13025 | -59.6095
0.13061 | -59.5725
0.13097 | -59.6005
0.13133 | -59.6994
0.1317 | -59.8628
0.13206 | -60.11
0.13242 | -60.4168
0.13278 | -60.7633
0.13314 | -61.1453
0.13351 | -61.5686
0.13387 | -62.0182
0.13423 | -62.4521
0.13459 | -62.8575
0.13495 | -63.2285
0.13531 | -63.5579
0.13567 | -63.5619
0.13604 | -63.7685
0.1364 | -63.8635
0.13676 | -63.8915
0.13712 | -63.8819
0.13748 | -63.8789
0.13785 | -63.846
0.13821 | -63.7891
0.13857 | -63.7674
0.13893 | -63.7802
0.13929 | -63.8113
0.13966 | -63.825
0.14002 | -63.8708
0.14038 | -63.9454
0.14074 | -64.009
0.1411 | -63.9938
0.14146 | -63.9342
0.14183 | -63.8345
0.14219 | -63.6965
0.14255 | -63.4739
0.14291 | -63.2011
0.14327 | -62.9475
0.14364 | -62.6875
0.144 | -62.404
0.14436 | -62.1487
0.14472 | -61.9296
0.14508 | -61.7483
0.14544 | -61.6299
0.14581 | -61.6852
0.14617 | -61.6706
0.14653 | -61.7007
0.14689 | -61.7811
0.14725 | -61.9228
0.14761 | -62.1107
0.14798 | -62.3417
0.14834 | -62.6187
0.1487 | -62.9315
0.14906 | -63.2782
0.14942 | -63.6477
0.14979 | -64.0074
0.15015 | -64.3965
0.15051 | -64.7638
0.15087 | -65.0566
0.15123 | -65.2997
0.1516 | -65.4832
0.15196 | -65.6133
0.15232 | -65.639
0.15268 | -65.6047
0.15304 | -65.5201
0.1534 | -65.4294
0.15377 | -65.3333
0.15413 | -65.1984
0.15449 | -65.0596
0.15485 | -64.9508
0.15521 | -64.8584
0.15557 | -64.7841
0.15594 | -64.735
0.1563 | -64.7682
0.15666 | -64.7464
0.15702 | -64.7135
0.15738 | -64.714
0.15774 | -64.726
0.15811 | -64.7569
0.15847 | -64.7861
0.15883 | -64.7993
0.15919 | -64.8208
0.15955 | -64.8518
0.15992 | -64.8495
0.16028 | -64.8392
0.16064 | -64.8382
0.161 | -64.8309
0.16136 | -64.8258
0.16173 | -64.7964
0.16209 | -64.7682
0.16245 | -64.738
0.16281 | -64.7101
0.16317 | -64.6889
0.16353 | -64.6829
0.1639 | -64.6747
0.16426 | -64.6475
0.16462 | -64.6414
0.16498 | -64.6344
0.16534 | -64.6263
0.1657 | -64.6408
0.16607 | -64.6816
0.16643 | -64.722
0.16679 | -64.7544
0.16715 | -64.7821
0.16751 | -64.7902
0.16788 | -64.7793
0.16824 | -64.7713
0.1686 | -64.7645
0.16896 | -64.7495
0.16932 | -64.7375
0.16969 | -64.6935
0.17005 | -64.6264
0.17041 | -64.5631
0.17077 | -64.5064
0.17113 | -64.4697
0.17149 | -64.4252
0.17186 | -64.3846
0.17222 | -64.3674
0.17258 | -64.3542
0.17294 | -64.3377
0.1733 | -64.3359
0.17367 | -64.357
0.17403 | -64.3906
0.17439 | -64.426
0.17475 | -64.4648
0.17511 | -64.5215
0.17547 | -64.5821
0.17583 | -64.63
0.1762 | -64.6831
0.17656 | -64.7272
0.17692 | -64.7488
0.17728 | -64.744
0.17764 | -64.7645
0.17801 | -64.7524
0.17837 | -64.7187
0.17873 | -64.6734
0.17909 | -64.6337
0.17945 | -64.5982
0.17982 | -64.5568
0.18018 | -64.5212
0.18054 | -64.4803
0.1809 | -64.4395
0.18126 | -64.403
0.18162 | -64.3719
0.18199 | -64.3506
0.18235 | -64.3239
0.18271 | -64.3038
0.18307 | -64.3076
0.18343 | -64.3163
0.18379 | -64.3356
0.18416 | -64.3517
0.18452 | -64.3849
0.18488 | -64.4304
0.18524 | -64.4799
0.1856 | -64.5321
0.18597 | -64.5733
0.18633 | -64.5985
0.18669 | -64.6295
0.18705 | -64.6877
0.18741 | -64.7696
0.18777 | -64.871
0.18814 | -64.9656
0.1885 | -65.0528
0.18886 | -65.1522
0.18922 | -65.2684
0.18958 | -65.3901
0.18995 | -65.5025
0.19031 | -65.6176
0.19067 | -65.7284
0.19103 | -65.8485
0.19139 | -65.9596
0.19176 | -66.0935
0.19212 | -66.2289
0.19248 | -66.3584
0.19284 | -66.4794
0.1932 | -66.6003
0.19356 | -66.7148
0.19392 | -66.8176
0.19429 | -66.9571
0.19465 | -67.0829
0.19501 | -67.1769
0.19537 | -67.2458
0.19573 | -67.3302
0.1961 | -67.4353
0.19646 | -67.4906
0.19682 | -67.4794
0.19718 | -67.4802
0.19754 | -67.4795
0.1979 | -67.4808
0.19827 | -67.445
0.19863 | -67.4232
0.19899 | -67.4026
0.19935 | -67.3562
0.19971 | -67.2969
0.20008 | -67.258
0.20044 | -67.2201
0.2008 | -67.1556
0.20116 | -67.068
0.20152 | -66.9671
0.20189 | -66.8436
0.20225 | -66.7324
0.20261 | -66.6337
0.20297 | -66.5428
0.20333 | -66.4503
0.20369 | -66.3481
0.20406 | -66.2249
0.20442 | -66.0828
0.20478 | -65.9406
0.20514 | -65.8142
0.2055 | -65.7004
0.20586 | -65.6028
0.20623 | -65.4984
0.20659 | -65.3905
0.20695 | -65.2961
0.20731 | -65.2099
0.20767 | -65.1274
0.20804 | -65.0593
0.2084 | -65.0102
0.20876 | -65.0079
0.20912 | -64.981
0.20948 | -64.971
0.20985 | -64.9667
0.21021 | -64.9938
0.21057 | -65.0289
0.21093 | -65.0789
0.21129 | -65.1547
0.21165 | -65.2554
0.21202 | -65.358
0.21238 | -65.4657
0.21274 | -65.5886
0.2131 | -65.7382
0.21346 | -65.9077
0.21382 | -66.0853
0.21419 | -66.2768
0.21455 | -66.4988
0.21491 | -66.7363
0.21527 | -66.973
0.21563 | -67.2027
0.21599 | -67.4431
0.21636 | -67.6747
0.21672 | -67.9168
0.21708 | -68.181
0.21744 | -68.4007
0.2178 | -68.5873
0.21817 | -68.78
0.21853 | -68.9653
0.21889 | -69.051
0.21925 | -69.2268
0.21961 | -69.415
0.21998 | -69.6073
0.22034 | -69.7562
0.2207 | -69.8304
0.22106 | -69.8677
0.22142 | -69.8984
0.22178 | -69.9434
0.22215 | -69.9767
0.22251 | -69.9838
0.22287 | -69.9727
0.22323 | -69.9649
0.22359 | -69.9511
0.22395 | -69.9288
0.22432 | -69.9157
0.22468 | -69.866
0.22504 | -69.79
0.2254 | -69.7321
0.22576 | -69.6579
0.22613 | -69.5658
0.22649 | -69.4763
0.22685 | -69.4027
0.22721 | -69.3241
0.22757 | -69.2438
0.22793 | -69.1719
0.2283 | -69.1093
0.22866 | -69.0317
0.22902 | -68.9581
0.22938 | -68.8997
0.22974 | -68.8229
0.23011 | -68.7608
0.23047 | -68.7042
0.23083 | -68.6478
0.23119 | -68.587
0.23155 | -68.5339
0.23191 | -68.4965
0.23228 | -68.4487
0.23264 | -68.408
0.233 | -68.373
0.23336 | -68.3472
0.23372 | -68.3172
0.23408 | -68.2895
0.23445 | -68.2719
0.23481 | -68.2641
0.23517 | -68.2659
0.23553 | -68.276
0.23589 | -68.2848
0.23626 | -68.297
0.23662 | -68.3191
0.23698 | -68.3502
0.23734 | -68.3941
0.2377 | -68.4464
0.23806 | -68.4946
0.23843 | -68.5371
0.23879 | -68.582
0.23915 | -68.6091
0.23951 | -68.6463
0.23987 | -68.6963
0.24024 | -68.7292
0.2406 | -68.7368
0.24096 | -68.7622
0.24132 | -68.7867
0.24168 | -68.8098
0.24205 | -68.8301
0.24241 | -68.8372
0.24277 | -68.8556
0.24313 | -68.8691
0.24349 | -68.8707
0.24385 | -68.8534
0.24422 | -68.8506
0.24458 | -68.8531
0.24494 | -68.8454
0.2453 | -68.849
0.24566 | -68.8688
0.24602 | -68.8731
0.24639 | -68.88
0.24675 | -68.9078
0.24711 | -68.946
0.24747 | -68.9783
0.24783 | -69.0153
0.2482 | -69.0668
0.24856 | -69.1422
0.24892 | -69.2296
0.24928 | -69.3245
0.24964 | -69.4281
0.25 | -69.5405
0.25037 | -69.6783
0.25073 | -69.7798
0.25109 | -69.894
0.25145 | -70.0429
0.25181 | -70.1847
0.25217 | -70.2808
0.25254 | -70.3714
0.2529 | -70.4691
0.25326 | -70.5726
0.25362 | -70.6585
0.25398 | -70.7408
0.25435 | -70.8144
0.25471 | -70.8784
0.25507 | -70.9184
0.25543 | -70.9448
0.25579 | -70.9573
0.25616 | -70.9478
0.25652 | -70.9358
0.25688 | -70.949
0.25724 | -70.932
0.2576 | -70.8466
0.25796 | -70.7459
0.25833 | -70.6789
0.25869 | -70.631
0.25905 | -70.533
0.25941 | -70.4335
0.25977 | -70.3409
0.26014 | -70.2292
0.2605 | -70.0988
0.26086 | -69.9788
0.26122 | -69.8683
0.26158 | -69.7394
0.26194 | -69.6193
0.26231 | -69.5171
0.26267 | -69.4015
0.26303 | -69.2821
0.26339 | -69.1823
0.26375 | -69.0862
0.26411 | -68.9952
0.26448 | -68.9151
0.26484 | -68.838
0.2652 | -68.7559
0.26556 | -68.6892
0.26592 | -68.6391
0.26629 | -68.5909
0.26665 | -68.5411
0.26701 | -68.4915
0.26737 | -68.4553
0.26773 | -68.4351
0.2681 | -68.4202
0.26846 | -68.4104
0.26882 | -68.4116
0.26918 | -68.4239
0.26954 | -68.4421
0.2699 | -68.4545
0.27026 | -68.4782
0.27063 | -68.5144
0.27099 | -68.5642
0.27135 | -68.6286
0.27171 | -68.6823
0.27207 | -68.7292
0.27244 | -68.776
0.2728 | -68.8288
0.27316 | -68.8854
0.27352 | -68.9398
0.27388 | -68.9954
0.27425 | -69.0596
0.27461 | -69.1212
0.27497 | -69.1839
0.27533 | -69.2579
0.27569 | -69.3137
0.27605 | -69.3636
0.27642 | -69.4207
0.27678 | -69.4969
0.27714 | -69.5789
0.2775 | -69.6437
0.27786 | -69.7037
0.27823 | -69.7704
0.27859 | -69.8438
0.27895 | -69.9165
0.27931 | -70.0062
0.27967 | -70.106
0.28003 | -70.1933
0.2804 | -70.2652
0.28076 | -70.3367
0.28112 | -70.4317
0.28148 | -70.5221
0.28184 | -70.6037
0.2822 | -70.6801
0.28257 | -70.7726
0.28293 | -70.865
0.28329 | -70.9376
0.28365 | -71.0059
0.28401 | -71.0715
0.28438 | -71.1417
0.28474 | -71.2101
0.2851 | -71.2869
0.28546 | -71.3429
0.28582 | -71.3866
0.28619 | -71.4214
0.28655 | -71.437
0.28691 | -71.446
0.28727 | -71.4484
0.28763 | -71.4338
0.28799 | -71.4286
0.28835 | -71.4294
0.28872 | -71.4203
0.28908 | -71.4
0.28944 | -71.3811
0.2898 | -71.361
0.29016 | -71.3207
0.29053 | -71.269
0.29089 | -71.2095
0.29125 | -71.1468
0.29161 | -71.084
0.29197 | -71.0057
0.29234 | -70.9446
0.2927 | -70.8891
0.29306 | -70.8331
0.29342 | -70.7696
0.29378 | -70.699
0.29414 | -70.6319
0.29451 | -70.5601
0.29487 | -70.4868
0.29523 | -70.4315
0.29559 | -70.3829
0.29595 | -70.3237
0.29631 | -70.2858
0.29668 | -70.2491
0.29704 | -70.2013
0.2974 | -70.1686
0.29776 | -70.1374
0.29812 | -70.1177
0.29849 | -70.101
0.29885 | -70.0825
0.29921 | -70.0682
0.29957 | -70.0601
0.29993 | -70.0442
0.30029 | -70.0364
0.30066 | -70.0284
0.30102 | -70.0351
0.30138 | -70.0497
0.30174 | -70.0546
0.3021 | -70.055
0.30247 | -70.0487
0.30283 | -70.0501
0.30319 | -70.052
0.30355 | -70.0575
0.30391 | -70.063
0.30428 | -70.0766
0.30464 | -70.0899
0.305 | -70.0962
0.30536 | -70.1023
0.30572 | -70.109
0.30608 | -70.1042
0.30645 | -70.0994
0.30681 | -70.1153
0.30717 | -70.1381
0.30753 | -70.1606
0.30789 | -70.2001
0.30825 | -70.2346
0.30862 | -70.246
0.30898 | -70.2468
0.30934 | -70.2494
0.3097 | -70.2451
0.31006 | -70.2319
0.31043 | -70.2213
0.31079 | -70.2124
0.31115 | -70.1972
0.31151 | -70.1791
0.31187 | -70.1576
0.31223 | -70.1306
0.3126 | -70.1064
0.31296 | -70.0587
0.31332 | -70.0277
0.31368 | -70.002
0.31404 | -69.966
0.3144 | -69.9248
0.31477 | -69.8773
0.31513 | -69.8411
0.31549 | -69.8031
0.31585 | -69.7753
0.31621 | -69.7552
0.31658 | -69.7394
0.31694 | -69.7233
0.3173 | -69.706
0.31766 | -69.698
0.31802 | -69.703
0.31838 | -69.7095
0.31875 | -69.7237
0.31911 | -69.7541
0.31947 | -69.7827
0.31983 | -69.8207
0.32019 | -69.8626
0.32056 | -69.9174
0.32092 | -69.9881
0.32128 | -70.0536
0.32164 | -70.1296
0.322 | -70.2302
0.32237 | -70.3274
0.32273 | -70.4283
0.32309 | -70.5365
0.32345 | -70.6575
0.32381 | -70.7607
0.32417 | -70.8801
0.32454 | -71.0027
0.3249 | -71.1242
0.32526 | -71.2492
0.32562 | -71.3722
0.32598 | -71.4906
0.32634 | -71.6026
0.32671 | -71.7157
0.32707 | -71.8298
0.32743 | -71.924
0.32779 | -72.0301
0.32815 | -72.143
0.32852 | -72.2369
0.32888 | -72.3301
0.32924 | -72.3949
0.3296 | -72.4358
0.32996 | -72.4615
0.33032 | -72.4794
0.33069 | -72.4956
0.33105 | -72.5036
0.33141 | -72.5137
0.33177 | -72.4933
0.33213 | -72.4512
0.33249 | -72.4114
0.33286 | -72.3594
0.33322 | -72.2819
0.33358 | -72.2198
0.33394 | -72.1288
0.3343 | -72.0566
0.33467 | -71.9845
0.33503 | -71.9174
0.33539 | -71.8523
0.33575 | -71.767
0.33611 | -71.6862
0.33648 | -71.6142
0.33684 | -71.5356
0.3372 | -71.4626
0.33756 | -71.399
0.33792 | -71.3429
0.33828 | -71.2853
0.33865 | -71.2174
0.33901 | -71.1551
0.33937 | -71.0863
0.33973 | -71.0275
0.34009 | -70.9794
0.34046 | -70.9374
0.34082 | -70.8984
0.34118 | -70.8602
0.34154 | -70.8306
0.3419 | -70.8091
0.34226 | -70.7935
0.34263 | -70.7847
0.34299 | -70.7834
0.34335 | -70.7817
0.34371 | -70.7755
0.34407 | -70.7789
0.34443 | -70.8222
0.3448 | -70.8376
0.34516 | -70.8717
0.34552 | -70.9081
0.34588 | -70.9425
0.34624 | -70.9808
0.3466 | -71.0295
0.34697 | -71.0785
0.34733 | -71.1302
0.34769 | -71.1952
0.34805 | -71.2594
0.34841 | -71.3134
0.34878 | -71.372
0.34914 | -71.4434
0.3495 | -71.5091
0.34986 | -71.5701
0.35022 | -71.633
0.35058 | -71.7071
0.35095 | -71.7778
0.35131 | -71.8396
0.35167 | -71.9072
0.35203 | -71.9631
0.35239 | -72.0136
0.35276 | -72.0645
0.35312 | -72.1189
0.35348 | -72.1867
0.35384 | -72.2351
0.3542 | -72.2811
0.35457 | -72.3143
0.35493 | -72.3682
0.35529 | -72.4034
0.35565 | -72.4168
0.35601 | -72.4402
0.35637 | -72.4887
0.35674 | -72.511
0.3571 | -72.5209
0.35746 | -72.5509
0.35782 | -72.5853
0.35818 | -72.5749
0.35855 | -72.5578
0.35891 | -72.5556
0.35927 | -72.5456
0.35963 | -72.5102
0.35999 | -72.4726
0.36035 | -72.438
0.36072 | -72.4109
0.36108 | -72.3758
0.36144 | -72.3312
0.3618 | -72.2909
0.36216 | -72.2407
0.36252 | -72.183
0.36289 | -72.1326
0.36325 | -72.0781
0.36361 | -72.0149
0.36397 | -71.9643
0.36433 | -71.9165
0.36469 | -71.8573
0.36506 | -71.8079
0.36542 | -71.7514
0.36578 | -71.6954
0.36614 | -71.6442
0.3665 | -71.6023
0.36687 | -71.5577
0.36723 | -71.4957
0.36759 | -71.4264
0.36795 | -71.3704
0.36831 | -71.3314
0.36867 | -71.3028
0.36904 | -71.2926
0.3694 | -71.2775
0.36976 | -71.2552
0.37012 | -71.239
0.37048 | -71.2323
0.37085 | -71.2214
0.37121 | -71.2163
0.37157 | -71.2184
0.37193 | -71.227
0.37229 | -71.2369
0.37266 | -71.2508
0.37302 | -71.2727
0.37338 | -71.2974
0.37374 | -71.3223
0.3741 | -71.3458
0.37446 | -71.3749
0.37483 | -71.4035
0.37519 | -71.435
0.37555 | -71.4802
0.37591 | -71.5111
0.37627 | -71.5451
0.37663 | -71.578
0.377 | -71.6139
0.37736 | -71.6564
0.37772 | -71.7094
0.37808 | -71.7641
0.37844 | -71.8146
0.37881 | -71.8629
0.37917 | -71.9133
0.37953 | -71.9617
0.37989 | -72.0188
0.38025 | -72.0797
0.38061 | -72.143
0.38098 | -72.2026
0.38134 | -72.2654
0.3817 | -72.3393
0.38206 | -72.4273
0.38242 | -72.5067
0.38279 | -72.584
0.38315 | -72.6632
0.38351 | -72.7275
0.38387 | -72.7766
0.38423 | -72.8411
0.3846 | -72.9237
0.38496 | -73.0005
0.38532 | -73.0669
0.38568 | -73.1419
0.38604 | -73.1871
0.3864 | -73.2641
0.38677 | -73.3235
0.38713 | -73.3646
0.38749 | -73.4101
0.38785 | -73.4551
0.38821 | -73.4933
0.38857 | -73.5156
0.38894 | -73.545
0.3893 | -73.5665
0.38966 | -73.5713
0.39002 | -73.5886
0.39038 | -73.6081
0.39075 | -73.6048
0.39111 | -73.5998
0.39147 | -73.5871
0.39183 | -73.5701
0.39219 | -73.5408
0.39255 | -73.5052
0.39292 | -73.4811
0.39328 | -73.4555
0.39364 | -73.4246
0.394 | -73.3985
0.39436 | -73.3952
0.39472 | -73.3984
0.39509 | -73.3747
0.39545 | -73.3296
0.39581 | -73.2914
0.39617 | -73.2347
0.39653 | -73.177
0.3969 | -73.1392
0.39726 | -73.1113
0.39762 | -73.0858
0.39798 | -73.0534
0.39834 | -73.0255
0.3987 | -72.9986
0.39907 | -72.975
0.39943 | -72.9588
0.39979 | -72.9441
0.40015 | -72.9177
0.40051 | -72.8942
0.40088 | -72.8828
0.40124 | -72.8784
0.4016 | -72.8743
0.40196 | -72.8613
0.40232 | -72.8487
0.40269 | -72.8313
0.40305 | -72.8162
0.40341 | -72.8063
0.40377 | -72.8023
0.40413 | -72.7907
0.40449 | -72.7723
0.40486 | -72.7523
0.40522 | -72.7234
0.40558 | -72.693
0.40594 | -72.6681
0.4063 | -72.6453
0.40666 | -72.6214
0.40703 | -72.5782
0.40739 | -72.5542
0.40775 | -72.53
0.40811 | -72.5004
0.40847 | -72.4678
0.40884 | -72.4377
0.4092 | -72.4095
0.40956 | -72.3805
0.40992 | -72.3506
0.41028 | -72.3236
0.41064 | -72.2864
0.41101 | -72.2443
0.41137 | -72.2157
0.41173 | -72.1865
0.41209 | -72.1573
0.41245 | -72.1252
0.41281 | -72.0937
0.41318 | -72.0652
0.41354 | -72.0279
0.4139 | -71.9883
0.41426 | -71.9595
0.41462 | -71.9367
0.41499 | -71.9111
0.41535 | -71.8909
0.41571 | -71.8814
0.41607 | -71.8655
0.41643 | -71.848
0.4168 | -71.8384
0.41716 | -71.8353
0.41752 | -71.8523
0.41788 | -71.8495
0.41824 | -71.853
0.4186 | -71.8599
0.41897 | -71.8469
0.41933 | -71.8354
0.41969 | -71.8382
0.42005 | -71.8461
0.42041 | -71.8582
0.42078 | -71.8684
0.42114 | -71.8879
0.4215 | -71.9147
0.42186 | -71.9337
0.42222 | -71.9542
0.42258 | -71.9748
0.42295 | -71.9888
0.42331 | -72.0112
0.42367 | -72.0385
0.42403 | -72.0696
0.42439 | -72.0954
0.42475 | -72.1165
0.42512 | -72.1451
0.42548 | -72.1802
0.42584 | -72.2125
0.4262 | -72.235
0.42656 | -72.2552
0.42692 | -72.2767
0.42729 | -72.301
0.42765 | -72.3267
0.42801 | -72.3526
0.42837 | -72.3733
0.42873 | -72.392
0.4291 | -72.4073
0.42946 | -72.425
0.42982 | -72.457
0.43018 | -72.4853
0.43054 | -72.4997
0.4309 | -72.5068
0.43127 | -72.52
0.43163 | -72.5411
0.43199 | -72.5608
0.43235 | -72.5829
0.43271 | -72.6043
0.43308 | -72.6241
0.43344 | -72.641
0.4338 | -72.6648
0.43416 | -72.6905
0.43452 | -72.7135
0.43489 | -72.7338
0.43525 | -72.7513
0.43561 | -72.7596
0.43597 | -72.7578
0.43633 | -72.7613
0.43669 | -72.7707
0.43706 | -72.7879
0.43742 | -72.8024
0.43778 | -72.8141
0.43814 | -72.8297
0.4385 | -72.839
0.43887 | -72.8437
0.43923 | -72.8491
0.43959 | -72.8569
0.43995 | -72.8634
0.44031 | -72.8709
0.44067 | -72.8806
0.44104 | -72.8949
0.4414 | -72.9105
0.44176 | -72.9214
0.44212 | -72.9362
0.44248 | -72.9474
0.44284 | -72.9496
0.44321 | -72.9515
0.44357 | -72.9551
0.44393 | -72.9682
0.44429 | -72.9844
0.44465 | -72.9946
0.44501 | -73.0018
0.44538 | -73.0069
0.44574 | -73.014
0.4461 | -73.0198
0.44646 | -73.0255
0.44682 | -73.0333
0.44719 | -73.0401
0.44755 | -73.0515
0.44791 | -73.0634
0.44827 | -73.0792
0.44863 | -73.1077
0.44899 | -73.1268
0.44936 | -73.1503
0.44972 | -73.1792
0.45008 | -73.2256
0.45044 | -73.2617
0.4508 | -73.2823
0.45117 | -73.3044
0.45153 | -73.3325
0.45189 | -73.3681
0.45225 | -73.4006
0.45261 | -73.4312
0.45298 | -73.4598
0.45334 | -73.4872
0.4537 | -73.5236
0.45406 | -73.5714
0.45442 | -73.6073
0.45478 | -73.641
0.45515 | -73.6711
0.45551 | -73.7068
0.45587 | -73.7403
0.45623 | -73.7791
0.45659 | -73.8189
0.45696 | -73.85
0.45732 | -73.8766
0.45768 | -73.9045
0.45804 | -73.9355
0.4584 | -73.961
0.45876 | -73.9849
0.45913 | -74.0011
0.45949 | -74.0257
0.45985 | -74.043
0.46021 | -74.0684
0.46057 | -74.0931
0.46093 | -74.1141
0.4613 | -74.1207
0.46166 | -74.1156
0.46202 | -74.1041
0.46238 | -74.0955
0.46274 | -74.0895
0.4631 | -74.0737
0.46347 | -74.0472
0.46383 | -74.0225
0.46419 | -73.9948
0.46455 | -73.967
0.46491 | -73.9377
0.46528 | -73.9063
0.46564 | -73.8735
0.466 | -73.8355
0.46636 | -73.7938
0.46672 | -73.7637
0.46708 | -73.7329
0.46745 | -73.7006
0.46781 | -73.6746
0.46817 | -73.6434
0.46853 | -73.6051
0.46889 | -73.5515
0.46926 | -73.4804
0.46962 | -73.3911
0.46998 | -73.3355
0.47034 | -73.2814
0.4707 | -73.2293
0.47107 | -73.1779
0.47143 | -73.1348
0.47179 | -73.0883
0.47215 | -73.041
0.47251 | -72.9945
0.47287 | -72.9428
0.47324 | -72.8886
0.4736 | -72.8404
0.47396 | -72.7889
0.47432 | -72.7268
0.47468 | -72.6726
0.47504 | -72.6259
0.47541 | -72.5775
0.47577 | -72.521
0.47613 | -72.4676
0.47649 | -72.4218
0.47685 | -72.3779
0.47722 | -72.3341
0.47758 | -72.2959
0.47794 | -72.2645
0.4783 | -72.2297
0.47866 | -72.1914
0.47902 | -72.161
0.47939 | -72.1348
0.47975 | -72.1031
0.48011 | -72.0883
0.48047 | -72.0618
0.48083 | -72.0368
0.48119 | -72.0147
0.48156 | -71.995
0.48192 | -71.9805
0.48228 | -71.9661
0.48264 | -71.9524
0.483 | -71.9421
0.48337 | -71.9344
0.48373 | -71.9308
0.48409 | -71.932
0.48445 | -71.9327
0.48481 | -71.9321
0.48518 | -71.9312
0.48554 | -71.9335
0.4859 | -71.9385
0.48626 | -71.947
0.48662 | -71.9572
0.48698 | -71.969
0.48735 | -71.9825
0.48771 | -72.0002
0.48807 | -72.0193
0.48843 | -72.0424
0.48879 | -72.0659
0.48916 | -72.0884
0.48952 | -72.1114
0.48988 | -72.1366
0.49024 | -72.1637
0.4906 | -72.2085
0.49096 | -72.2407
0.49133 | -72.2704
0.49169 | -72.2926
0.49205 | -72.3253
0.49241 | -72.3639
0.49277 | -72.3951
0.49313 | -72.4298
0.4935 | -72.4714
0.49386 | -72.5204
0.49422 | -72.5672
0.49458 | -72.6137
0.49494 | -72.6648
0.49531 | -72.7245
0.49567 | -72.7877
0.49603 | -72.8513
0.49639 | -72.9143
0.49675 | -72.9765
0.49711 | -73.0359
0.49748 | -73.0906
0.49784 | -73.1467
0.4982 | -73.2031
0.49856 | -73.2563
0.49892 | -73.3082
0.49928 | -73.3625
0.49965 | -73.4087
0.50001 | -73.464
0.50037 | -73.5237
0.50073 | -73.5661
0.50109 | -73.6144
0.50146 | -73.6605
0.50182 | -73.7063
0.50218 | -73.7506
0.50254 | -73.7902
0.5029 | -73.8242
0.50326 | -73.857
0.50363 | -73.9105
0.50399 | -73.9674
0.50435 | -74.0117
0.50471 | -74.0432
0.50507 | -74.0729
0.50544 | -74.1054
0.5058 | -74.1329
0.50616 | -74.148
0.50652 | -74.1559
0.50688 | -74.1734
0.50724 | -74.1904
0.50761 | -74.2109
0.50797 | -74.238
0.50833 | -74.2608
0.50869 | -74.2825
0.50905 | -74.3104
0.50942 | -74.3309
0.50978 | -74.3463
0.51014 | -74.3626
0.5105 | -74.3891
0.51086 | -74.4137
0.51123 | -74.4308
0.51159 | -74.4461
0.51195 | -74.4626
0.51231 | -74.4801
0.51267 | -74.4915
0.51303 | -74.5073
0.5134 | -74.5306
0.51376 | -74.5577
0.51412 | -74.5758
0.51448 | -74.5905
0.51484 | -74.614
0.51521 | -74.6388
0.51557 | -74.6548
0.51593 | -74.6787
0.51629 | -74.7066
0.51665 | -74.7329
0.51701 | -74.7573
0.51738 | -74.7799
0.51774 | -74.8054
0.5181 | -74.8316
0.51846 | -74.8558
0.51882 | -74.8869
0.51919 | -74.9234
0.51955 | -74.9543
0.51991 | -74.9891
0.52027 | -75.0292
0.52063 | -75.0754
0.52099 | -75.1212
0.52136 | -75.1668
0.52172 | -75.2338
0.52208 | -75.2875
0.52244 | -75.344
0.5228 | -75.3994
0.52316 | -75.4565
0.52353 | -75.513
0.52389 | -75.5721
0.52425 | -75.6347
0.52461 | -75.7052
0.52497 | -75.7844
0.52533 | -75.8583
0.5257 | -75.9314
0.52606 | -76.008
0.52642 | -76.0911
0.52678 | -76.1733
0.52714 | -76.2581
0.52751 | -76.3572
0.52787 | -76.4724
0.52823 | -76.5707
0.52859 | -76.6632
0.52895 | -76.7655
0.52931 | -76.8578
0.52968 | -76.9349
0.53004 | -77.0027
0.5304 | -77.0709
0.53076 | -77.1581
0.53112 | -77.2417
0.53149 | -77.3138
0.53185 | -77.391
0.53221 | -77.4099
0.53257 | -77.4668
0.53293 | -77.5355
0.53329 | -77.6039
0.53366 | -77.6709
0.53402 | -77.7426
0.53438 | -77.797
0.53474 | -77.8342
0.5351 | -77.8579
0.53547 | -77.8841
0.53583 | -77.9022
0.53619 | -77.9317
0.53655 | -77.9442
0.53691 | -77.9379
0.53727 | -77.9285
0.53764 | -77.9253
0.538 | -77.9122
0.53836 | -77.8773
0.53872 | -77.8498
0.53908 | -77.8223
0.53944 | -77.7942
0.53981 | -77.7652
0.54017 | -77.7367
0.54053 | -77.6857
0.54089 | -77.6158
0.54125 | -77.5464
0.54162 | -77.4714
0.54198 | -77.3834
0.54234 | -77.2918
0.5427 | -77.182
0.54306 | -77.0983
0.54343 | -77.0326
0.54379 | -76.9692
0.54415 | -76.8785
0.54451 | -76.7957
0.54487 | -76.7151
0.54523 | -76.6318
0.5456 | -76.5495
0.54596 | -76.4771
0.54632 | -76.4079
0.54668 | -76.3321
0.54704 | -76.2517
0.54741 | -76.1711
0.54777 | -76.1005
0.54813 | -76.0308
0.54849 | -75.9613
0.54885 | -75.8919
0.54921 | -75.8206
0.54958 | -75.7508
0.54994 | -75.6882
0.5503 | -75.6333
0.55066 | -75.5792
0.55102 | -75.5237
0.55139 | -75.4736
0.55175 | -75.4287
0.55211 | -75.3877
0.55247 | -75.3489
0.55283 | -75.3067
0.55319 | -75.2846
0.55356 | -75.247
0.55392 | -75.2094
0.55428 | -75.179
0.55464 | -75.1493
0.555 | -75.1179
0.55536 | -75.0943
0.55573 | -75.0744
0.55609 | -75.0583
0.55645 | -75.0461
0.55681 | -75.0336
0.55717 | -75.0233
0.55754 | -75.0075
0.5579 | -74.9976
0.55826 | -74.9921
0.55862 | -74.9874
0.55898 | -74.9855
0.55934 | -74.9866
0.55971 | -74.9868
0.56007 | -74.9863
0.56043 | -74.9909
0.56079 | -74.9968
0.56115 | -75.0058
0.56151 | -75.0188
0.56188 | -75.0325
0.56224 | -75.0479
0.5626 | -75.0663
0.56296 | -75.0855
0.56332 | -75.1024
0.56369 | -75.1384
0.56405 | -75.1556
0.56441 | -75.1694
0.56477 | -75.1873
0.56513 | -75.2085
0.56549 | -75.2313
0.56586 | -75.254
0.56622 | -75.2797
0.56658 | -75.3073
0.56694 | -75.3353
0.5673 | -75.3656
0.56767 | -75.4006
0.56803 | -75.4388
0.56839 | -75.4682
0.56875 | -75.4977
0.56911 | -75.5298
0.56947 | -75.5576
0.56984 | -75.587
0.5702 | -75.6187
0.57056 | -75.6485
0.57092 | -75.6816
0.57128 | -75.7256
0.57165 | -75.7643
0.57201 | -75.7908
0.57237 | -75.8125
0.57273 | -75.8432
0.57309 | -75.8714
0.57346 | -75.897
0.57382 | -75.9252
0.57418 | -75.9544
0.57454 | -75.9832
0.5749 | -75.9973
0.57526 | -76.0211
0.57563 | -76.0495
0.57599 | -76.0759
0.57635 | -76.1008
0.57671 | -76.1249
0.57707 | -76.143
0.57743 | -76.1641
0.5778 | -76.1965
0.57816 | -76.2207
0.57852 | -76.237
0.57888 | -76.2511
0.57924 | -76.2693
0.57961 | -76.2899
0.57997 | -76.3109
0.58033 | -76.3321
0.58069 | -76.3558
0.58105 | -76.377
0.58141 | -76.4026
0.58178 | -76.4215
0.58214 | -76.4332
0.5825 | -76.4535
0.58286 | -76.4772
0.58322 | -76.4957
0.58359 | -76.5124
0.58395 | -76.5351
0.58431 | -76.5583
0.58467 | -76.5755
0.58503 | -76.5933
0.58539 | -76.6112
0.58576 | -76.6284
0.58612 | -76.647
0.58648 | -76.6651
0.58684 | -76.6778
0.5872 | -76.6937
0.58757 | -76.7184
0.58793 | -76.7351
0.58829 | -76.749
0.58865 | -76.7701
0.58901 | -76.7987
0.58937 | -76.8318
0.58974 | -76.8666
0.5901 | -76.8974
0.59046 | -76.9222
0.59082 | -76.9406
0.59118 | -76.9657
0.59154 | -76.9946
0.59191 | -77.0205
0.59227 | -77.0475
0.59263 | -77.0789
0.59299 | -77.1178
0.59335 | -77.1487
0.59372 | -77.1688
0.59408 | -77.2014
0.59444 | -77.2382
0.5948 | -77.2709
0.59516 | -77.3006
0.59552 | -77.3285
0.59589 | -77.3617
0.59625 | -77.3883
0.59661 | -77.4148
0.59697 | -77.4382
0.59733 | -77.4578
0.59769 | -77.4805
0.59806 | -77.5044
0.59842 | -77.5396
0.59878 | -77.5498
0.59914 | -77.5503
0.5995 | -77.5596
0.59987 | -77.5689
0.60023 | -77.5774
0.60059 | -77.5901
0.60095 | -77.6083
0.60131 | -77.6215
0.60167 | -77.6275
0.60204 | -77.6287
0.6024 | -77.6269
0.60276 | -77.622
0.60312 | -77.6131
0.60348 | -77.6053
0.60385 | -77.5974
0.60421 | -77.5892
0.60457 | -77.5707
0.60493 | -77.5523
0.60529 | -77.5215
0.60566 | -77.5157
0.60602 | -77.5136
0.60638 | -77.5038
0.60674 | -77.4923
0.6071 | -77.4765
0.60746 | -77.457
0.60783 | -77.4334
0.60819 | -77.4078
0.60855 | -77.3824
0.60891 | -77.3559
0.60927 | -77.3299
0.60964 | -77.3046
0.61 | -77.2845
0.61036 | -77.2651
0.61072 | -77.2416
0.61108 | -77.2107
0.61144 | -77.18
0.61181 | -77.1521
0.61217 | -77.1187
0.61253 | -77.0886
0.61289 | -77.0638
0.61325 | -77.0371
0.61362 | -77.0101
0.61398 | -76.9922
0.61434 | -76.9688
0.6147 | -76.9351
0.61506 | -76.9021
0.61542 | -76.8774
0.61579 | -76.8606
0.61615 | -76.8323
0.61651 | -76.8115
0.61687 | -76.7976
0.61723 | -76.79
0.61759 | -76.7715
0.61796 | -76.752
0.61832 | -76.7331
0.61868 | -76.7205
0.61904 | -76.7073
0.6194 | -76.6941
0.61977 | -76.6756
0.62013 | -76.66
0.62049 | -76.6495
0.62085 | -76.6412
0.62121 | -76.6289
0.62157 | -76.6109
0.62194 | -76.5994
0.6223 | -76.5897
0.62266 | -76.5811
0.62302 | -76.5736
0.62338 | -76.5688
0.62374 | -76.5676
0.62411 | -76.5703
0.62447 | -76.5675
0.62483 | -76.5629
0.62519 | -76.5597
0.62555 | -76.5539
0.62592 | -76.5501
0.62628 | -76.5543
0.62664 | -76.5532
0.627 | -76.5522
0.62736 | -76.5542
0.62772 | -76.5562
0.62809 | -76.5578
0.62845 | -76.5568
0.62881 | -76.5553
0.62917 | -76.5567
0.62953 | -76.5618
0.6299 | -76.5719
0.63026 | -76.5827
0.63062 | -76.5924
0.63098 | -76.6032
0.63134 | -76.6117
0.6317 | -76.6077
0.63207 | -76.6055
0.63243 | -76.6103
0.63279 | -76.6167
0.63315 | -76.623
0.63351 | -76.629
0.63387 | -76.6412
0.63424 | -76.6521
0.6346 | -76.6594
0.63496 | -76.6696
0.63532 | -76.6821
0.63568 | -76.6942
0.63605 | -76.7037
0.63641 | -76.713
0.63677 | -76.7206
0.63713 | -76.7353
0.63749 | -76.7445
0.63785 | -76.7518
0.63822 | -76.7533
0.63858 | -76.7535
0.63894 | -76.7568
0.6393 | -76.7671
0.63966 | -76.7781
0.64003 | -76.7845
0.64039 | -76.7905
0.64075 | -76.8031
0.64111 | -76.8197
0.64147 | -76.8337
0.64184 | -76.8444
0.6422 | -76.8525
0.64256 | -76.8678
0.64292 | -76.8874
0.64328 | -76.9025
0.64364 | -76.914
0.64401 | -76.9205
0.64437 | -76.9213
0.64473 | -76.9189
0.64509 | -76.9131
0.64545 | -76.9089
0.64582 | -76.9055
0.64618 | -76.9046
0.64654 | -76.9039
0.6469 | -76.8901
0.64726 | -76.8918
0.64762 | -76.8921
0.64799 | -76.8923
0.64835 | -76.8874
0.64871 | -76.8738
0.64907 | -76.861
0.64943 | -76.8528
0.6498 | -76.8422
0.65016 | -76.826
0.65052 | -76.8121
0.65088 | -76.8011
0.65124 | -76.7865
0.6516 | -76.7678
0.65197 | -76.7514
0.65233 | -76.7333
0.65269 | -76.7117
0.65305 | -76.6911
0.65341 | -76.6728
0.65377 | -76.6528
0.65414 | -76.6299
0.6545 | -76.6136
0.65486 | -76.5986
0.65522 | -76.581
0.65558 | -76.5598
0.65595 | -76.5358
0.65631 | -76.5132
0.65667 | -76.4925
0.65703 | -76.4672
0.65739 | -76.4371
0.65775 | -76.4066
0.65812 | -76.3719
0.65848 | -76.3425
0.65884 | -76.3183
0.6592 | -76.2952
0.65956 | -76.2781
0.65992 | -76.2605
0.66029 | -76.2345
0.66065 | -76.2046
0.66101 | -76.1801
0.66137 | -76.1591
0.66173 | -76.1371
0.6621 | -76.116
0.66246 | -76.0973
0.66282 | -76.0781
0.66318 | -76.0565
0.66354 | -76.0394
0.6639 | -76.024
0.66427 | -76.009
0.66463 | -75.9926
0.66499 | -75.9812
0.66535 | -75.97
0.66571 | -75.9531
0.66608 | -75.9342
0.66644 | -75.9175
0.6668 | -75.9138
0.66716 | -75.9097
0.66752 | -75.9037
0.66788 | -75.9162
0.66825 | -75.9142
0.66861 | -75.9111
0.66897 | -75.908
0.66933 | -75.9053
0.66969 | -75.9001
0.67006 | -75.8928
0.67042 | -75.887
0.67078 | -75.8865
0.67114 | -75.8901
0.6715 | -75.8946
0.67187 | -75.9002
0.67223 | -75.9073
0.67259 | -75.9128
0.67295 | -75.9173
0.67331 | -75.9241
0.67367 | -75.9296
0.67404 | -75.9378
0.6744 | -75.9491
0.67476 | -75.963
0.67512 | -75.971
0.67548 | -75.9746
0.67584 | -75.9834
0.67621 | -75.9933
0.67657 | -75.9992
0.67693 | -76.0031
0.67729 | -76.0121
0.67765 | -76.0298
0.67802 | -76.0409
0.67838 | -76.0463
0.67874 | -76.055
0.6791 | -76.063
0.67946 | -76.0722
0.67982 | -76.0811
0.68019 | -76.09
0.68055 | -76.0943
0.68091 | -76.0962
0.68127 | -76.1025
0.68163 | -76.1089
0.682 | -76.1102
0.68236 | -76.1093
0.68272 | -76.1112
0.68308 | -76.1182
0.68344 | -76.1273
0.6838 | -76.1344
0.68417 | -76.1399
0.68453 | -76.1454
0.68489 | -76.1545
0.68525 | -76.1654
0.68561 | -76.1729
0.68598 | -76.173
0.68634 | -76.176
0.6867 | -76.1835
0.68706 | -76.1898
0.68742 | -76.193
0.68778 | -76.1964
0.68815 | -76.2008
0.68851 | -76.2068
0.68887 | -76.2078
0.68923 | -76.2137
0.68959 | -76.2241
0.68995 | -76.2324
0.69032 | -76.2381
0.69068 | -76.2428
0.69104 | -76.2487
0.6914 | -76.2568
0.69176 | -76.2686
0.69212 | -76.2854
0.69249 | -76.3001
0.69285 | -76.3122
0.69321 | -76.3246
0.69357 | -76.3367
0.69393 | -76.3493
0.6943 | -76.3596
0.69466 | -76.3677
0.69502 | -76.374
0.69538 | -76.3805
0.69574 | -76.3945
0.6961 | -76.4187
0.69647 | -76.4365
0.69683 | -76.4471
0.69719 | -76.4588
0.69755 | -76.4733
0.69791 | -76.4867
0.69828 | -76.5004
0.69864 | -76.5178
0.699 | -76.5393
0.69936 | -76.5666
0.69972 | -76.5856
0.70008 | -76.604
0.70045 | -76.6232
0.70081 | -76.6455
0.70117 | -76.666
0.70153 | -76.6864
0.70189 | -76.7102
0.70226 | -76.7321
0.70262 | -76.752
0.70298 | -76.7702
0.70334 | -76.7897
0.7037 | -76.8119
0.70407 | -76.834
0.70443 | -76.8501
0.70479 | -76.8698
0.70515 | -76.8935
0.70551 | -76.9203
0.70587 | -76.9491
0.70624 | -76.9826
0.7066 | -77.0128
0.70696 | -77.0359
0.70732 | -77.0583
0.70768 | -77.0821
0.70805 | -77.1065
0.70841 | -77.1297
0.70877 | -77.1564
0.70913 | -77.1798
0.70949 | -77.1932
0.70985 | -77.2181
0.71022 | -77.2393
0.71058 | -77.2603
0.71094 | -77.2832
0.7113 | -77.3049
0.71166 | -77.3288
0.71202 | -77.348
0.71239 | -77.3639
0.71275 | -77.385
0.71311 | -77.413
0.71347 | -77.4403
0.71383 | -77.4631
0.7142 | -77.488
0.71456 | -77.5152
0.71492 | -77.5491
0.71528 | -77.5682
0.71564 | -77.5826
0.716 | -77.599
0.71637 | -77.6083
0.71673 | -77.6114
0.71709 | -77.6236
0.71745 | -77.6427
0.71781 | -77.6624
0.71818 | -77.6802
0.71854 | -77.7035
0.7189 | -77.7323
0.71926 | -77.7508
0.71962 | -77.7617
0.71998 | -77.77
0.72035 | -77.7831
0.72071 | -77.7932
0.72107 | -77.7959
0.72143 | -77.7989
0.72179 | -77.8096
0.72215 | -77.8233
0.72252 | -77.8333
0.72288 | -77.8444
0.72324 | -77.8532
0.7236 | -77.8621
0.72396 | -77.87
0.72433 | -77.8821
0.72469 | -77.8976
0.72505 | -77.9077
0.72541 | -77.9165
0.72577 | -77.9293
0.72613 | -77.9417
0.7265 | -77.9476
0.72686 | -77.9516
0.72722 | -77.9625
0.72758 | -77.9752
0.72794 | -77.9797
0.72831 | -77.9821
0.72867 | -77.9863
0.72903 | -77.9894
0.72939 | -77.9897
0.72975 | -77.9893
0.73011 | -77.9987
0.73048 | -77.995
0.73084 | -78
0.7312 | -77.9972
0.73156 | -77.996
0.73192 | -78.0027
0.73228 | -78.0166
0.73265 | -78.0287
0.73301 | -78.0357
0.73337 | -78.0406
0.73373 | -78.0426
0.73409 | -78.0417
0.73446 | -78.0407
0.73482 | -78.0406
0.73518 | -78.0375
0.73554 | -78.0307
0.7359 | -78.0369
0.73626 | -78.042
0.73663 | -78.0437
0.73699 | -78.0441
0.73735 | -78.0462
0.73771 | -78.0486
0.73807 | -78.0482
0.73844 | -78.0466
0.7388 | -78.0515
0.73916 | -78.0618
0.73952 | -78.0672
0.73988 | -78.07
0.74025 | -78.0728
0.74061 | -78.0727
0.74097 | -78.0778
0.74133 | -78.0834
0.74169 | -78.093
0.74205 | -78.1036
0.74242 | -78.1089
0.74278 | -78.1085
0.74314 | -78.1107
0.7435 | -78.1157
0.74386 | -78.1134
0.74423 | -78.1118
0.74459 | -78.1142
0.74495 | -78.1196
0.74531 | -78.1288
0.74567 | -78.1368
0.74603 | -78.1436
0.7464 | -78.1463
0.74676 | -78.1477
0.74712 | -78.1544
0.74748 | -78.1619
0.74784 | -78.1683
0.74821 | -78.173
0.74857 | -78.174
0.74893 | -78.1761
0.74929 | -78.1791
0.74965 | -78.1818
0.75001 | -78.1836
0.75038 | -78.1812
0.75074 | -78.1805
0.7511 | -78.1841
0.75146 | -78.1913
0.75182 | -78.1972
0.75218 | -78.1963
0.75255 | -78.1954
0.75291 | -78.1943
0.75327 | -78.1951
0.75363 | -78.1987
0.75399 | -78.2036
0.75435 | -78.2067
0.75472 | -78.2134
0.75508 | -78.2191
0.75544 | -78.2238
0.7558 | -78.2328
0.75616 | -78.2378
0.75653 | -78.2403
0.75689 | -78.2416
0.75725 | -78.2406
0.75761 | -78.2357
0.75797 | -78.2269
0.75833 | -78.2227
0.7587 | -78.2215
0.75906 | -78.2217
0.75942 | -78.2226
0.75978 | -78.2255
0.76014 | -78.2284
0.76051 | -78.234
0.76087 | -78.2423
0.76123 | -78.2468
0.76159 | -78.2483
0.76195 | -78.2506
0.76231 | -78.2525
0.76268 | -78.2558
0.76304 | -78.2589
0.7634 | -78.2585
0.76376 | -78.2588
0.76412 | -78.2619
0.76449 | -78.2668
0.76485 | -78.267
0.76521 | -78.2678
0.76557 | -78.2714
0.76593 | -78.2779
0.76629 | -78.2843
0.76666 | -78.2843
0.76702 | -78.2859
0.76738 | -78.2922
0.76774 | -78.2967
0.7681 | -78.3
0.76846 | -78.3132
0.76883 | -78.3305
0.76919 | -78.3361
0.76955 | -78.3335
0.76991 | -78.332
0.77027 | -78.3362
0.77064 | -78.3402
0.771 | -78.3434
0.77136 | -78.3533
0.77172 | -78.3627
0.77208 | -78.3669
0.77245 | -78.3694
0.77281 | -78.3729
0.77317 | -78.374
0.77353 | -78.3784
0.77389 | -78.389
0.77425 | -78.3969
0.77462 | -78.3982
0.77498 | -78.3974
0.77534 | -78.3984
0.7757 | -78.402
0.77606 | -78.4055
0.77643 | -78.4065
0.77679 | -78.4096
0.77715 | -78.4148
0.77751 | -78.4198
0.77787 | -78.4223
0.77823 | -78.4246
0.7786 | -78.4286
0.77896 | -78.4335
0.77932 | -78.4375
0.77968 | -78.443
0.78004 | -78.448
0.78041 | -78.4511
0.78077 | -78.4562
0.78113 | -78.4652
0.78149 | -78.4728
0.78185 | -78.4819
0.78221 | -78.4937
0.78258 | -78.4918
0.78294 | -78.4994
0.7833 | -78.5069
0.78366 | -78.5162
0.78402 | -78.5287
0.78438 | -78.543
0.78475 | -78.5559
0.78511 | -78.5644
0.78547 | -78.573
0.78583 | -78.5792
0.78619 | -78.5855
0.78656 | -78.5946
0.78692 | -78.6055
0.78728 | -78.6147
0.78764 | -78.6183
0.788 | -78.6226
0.78836 | -78.6308
0.78873 | -78.6427
0.78909 | -78.6576
0.78945 | -78.6639
0.78981 | -78.6685
0.79017 | -78.6779
0.79053 | -78.6891
0.7909 | -78.6971
0.79126 | -78.7069
0.79162 | -78.7211
0.79198 | -78.7392
0.79234 | -78.7557
0.79271 | -78.7699
0.79307 | -78.7799
0.79343 | -78.7897
0.79379 | -78.7953
0.79415 | -78.8001
0.79451 | -78.81
0.79488 | -78.8201
0.79524 | -78.8278
0.7956 | -78.8402
0.79596 | -78.8482
0.79632 | -78.8535
0.79669 | -78.8612
0.79705 | -78.8726
0.79741 | -78.8819
0.79777 | -78.8919
0.79813 | -78.9053
0.79849 | -78.9167
0.79886 | -78.9236
0.79922 | -78.9252
0.79958 | -78.92
0.79994 | -78.916
0.8003 | -78.9205
0.80067 | -78.9273
0.80103 | -78.933
0.80139 | -78.9363
0.80175 | -78.9381
0.80211 | -78.9373
0.80248 | -78.9353
0.80284 | -78.9364
0.8032 | -78.9382
0.80356 | -78.932
0.80392 | -78.925
0.80428 | -78.9201
0.80465 | -78.9183
0.80501 | -78.9193
0.80537 | -78.9195
0.80573 | -78.9159
0.80609 | -78.9076
0.80646 | -78.9016
0.80682 | -78.9006
0.80718 | -78.9007
0.80754 | -78.8994
0.8079 | -78.899
0.80826 | -78.8973
0.80863 | -78.8918
0.80899 | -78.8863
0.80935 | -78.8862
0.80971 | -78.8817
0.81007 | -78.8754
0.81043 | -78.8704
0.8108 | -78.8711
0.81116 | -78.8724
0.81152 | -78.8702
0.81188 | -78.8639
0.81224 | -78.8565
0.81261 | -78.8486
0.81297 | -78.8439
0.81333 | -78.8388
0.81369 | -78.8339
0.81405 | -78.8269
0.81441 | -78.823
0.81478 | -78.8192
0.81514 | -78.815
0.8155 | -78.8097
0.81586 | -78.8025
0.81622 | -78.7959
0.81659 | -78.7892
0.81695 | -78.7826
0.81731 | -78.7788
0.81767 | -78.7762
0.81803 | -78.7728
0.81839 | -78.7644
0.81876 | -78.7562
0.81912 | -78.7553
0.81948 | -78.7544
0.81984 | -78.7548
0.8202 | -78.7561
0.82056 | -78.755
0.82093 | -78.75
0.82129 | -78.7452
0.82165 | -78.7423
0.82201 | -78.7413
0.82237 | -78.7487
0.82274 | -78.7545
0.8231 | -78.7555
0.82346 | -78.7574
0.82382 | -78.7588
0.82418 | -78.7568
0.82454 | -78.7581
0.82491 | -78.7508
0.82527 | -78.7519
0.82563 | -78.7587
0.82599 | -78.7656
0.82635 | -78.768
0.82671 | -78.7708
0.82708 | -78.7767
0.82744 | -78.7774
0.8278 | -78.7723
0.82816 | -78.7677
0.82852 | -78.7727
0.82889 | -78.776
0.82925 | -78.7786
0.82961 | -78.7833
0.82997 | -78.7892
0.83033 | -78.7896
0.83069 | -78.7863
0.83106 | -78.7876
0.83142 | -78.7897
0.83178 | -78.7903
0.83214 | -78.7901
0.8325 | -78.7937
0.83287 | -78.7981
0.83323 | -78.7997
0.83359 | -78.7989
0.83395 | -78.7917
0.83431 | -78.785
0.83467 | -78.7855
0.83504 | -78.7812
0.8354 | -78.7717
0.83576 | -78.7676
0.83612 | -78.7729
0.83648 | -78.7728
0.83685 | -78.7634
0.83721 | -78.7613
0.83757 | -78.7599
0.83793 | -78.756
0.83829 | -78.7498
0.83866 | -78.7433
0.83902 | -78.7371
0.83938 | -78.7297
0.83974 | -78.7313
0.8401 | -78.731
0.84046 | -78.7266
0.84083 | -78.7201
0.84119 | -78.7103
0.84155 | -78.6998
0.84191 | -78.6919
0.84227 | -78.6856
0.84264 | -78.6767
0.843 | -78.6682
0.84336 | -78.6586
0.84372 | -78.6475
0.84408 | -78.6354
0.84444 | -78.6258
0.84481 | -78.6222
0.84517 | -78.6179
0.84553 | -78.6063
0.84589 | -78.5971
0.84625 | -78.5854
0.84661 | -78.5717
0.84698 | -78.5571
0.84734 | -78.5391
0.8477 | -78.5177
0.84806 | -78.4965
0.84842 | -78.479
0.84879 | -78.4635
0.84915 | -78.4485
0.84951 | -78.4313
0.84987 | -78.4111
0.85023 | -78.3911
0.85059 | -78.3759
0.85096 | -78.3633
0.85132 | -78.3514
0.85168 | -78.3405
0.85204 | -78.3274
0.8524 | -78.3128
0.85276 | -78.2941
0.85313 | -78.2802
0.85349 | -78.2644
0.85385 | -78.2481
0.85421 | -78.2327
0.85457 | -78.2184
0.85494 | -78.2039
0.8553 | -78.1917
0.85566 | -78.1786
0.85602 | -78.16
0.85638 | -78.1398
0.85674 | -78.1239
0.85711 | -78.1101
0.85747 | -78.098
0.85783 | -78.0915
0.85819 | -78.0854
0.85855 | -78.0739
0.85892 | -78.06
0.85928 | -78.0421
0.85964 | -78.0231
0.86 | -78.0094
0.86036 | -77.9991
0.86072 | -77.9881
0.86109 | -77.9733
0.86145 | -77.9592
0.86181 | -77.9485
0.86217 | -77.9395
0.86253 | -77.9343
0.86289 | -77.928
0.86326 | -77.9175
0.86362 | -77.9085
0.86398 | -77.9017
0.86434 | -77.8953
0.8647 | -77.8891
0.86507 | -77.8793
0.86543 | -77.8669
0.86579 | -77.8529
0.86615 | -77.841
0.86651 | -77.8312
0.86687 | -77.8253
0.86724 | -77.8201
0.8676 | -77.815
0.86796 | -77.8086
0.86832 | -77.8001
0.86868 | -77.7934
0.86905 | -77.7902
0.86941 | -77.7884
0.86977 | -77.784
0.87013 | -77.78
0.87049 | -77.7776
0.87086 | -77.777
0.87122 | -77.7762
0.87158 | -77.7783
0.87194 | -77.783
0.8723 | -77.7842
0.87266 | -77.7816
0.87303 | -77.7789
0.87339 | -77.7781
0.87375 | -77.7745
0.87411 | -77.7708
0.87447 | -77.7714
0.87484 | -77.773
0.8752 | -77.7748
0.87556 | -77.7774
0.87592 | -77.7788
0.87628 | -77.7801
0.87664 | -77.786
0.87701 | -77.7879
0.87737 | -77.7855
0.87773 | -77.7847
0.87809 | -77.7894
0.87845 | -77.7937
0.87882 | -77.7943
0.87918 | -77.7951
0.87954 | -77.7965
0.8799 | -77.7991
0.88026 | -77.8029
0.88062 | -77.8054
0.88099 | -77.8064
0.88135 | -77.8092
0.88171 | -77.8148
0.88207 | -77.8195
0.88243 | -77.8216
0.88279 | -77.8236
0.88316 | -77.826
0.88352 | -77.8308
0.88388 | -77.8371
0.88424 | -77.8449
0.8846 | -77.8535
0.88497 | -77.8651
0.88533 | -77.8709
0.88569 | -77.8742
0.88605 | -77.8784
0.88641 | -77.8852
0.88677 | -77.8931
0.88714 | -77.9052
0.8875 | -77.9144
0.88786 | -77.9265
0.88822 | -77.9423
0.88858 | -77.9591
0.88894 | -77.9697
0.88931 | -77.9773
0.88967 | -77.986
0.89003 | -77.9968
0.89039 | -78.0068
0.89075 | -78.0196
0.89112 | -78.0352
0.89148 | -78.0496
0.89184 | -78.0647
0.8922 | -78.0798
0.89256 | -78.0949
0.89292 | -78.109
0.89329 | -78.1236
0.89365 | -78.1386
0.89401 | -78.1553
0.89437 | -78.1762
0.89473 | -78.1967
0.8951 | -78.2067
0.89546 | -78.219
0.89582 | -78.2382
0.89618 | -78.2632
0.89654 | -78.2852
0.8969 | -78.3021
0.89727 | -78.3205
0.89763 | -78.3395
0.89799 | -78.3628
0.89835 | -78.3824
0.89871 | -78.4005
0.89908 | -78.4193
0.89944 | -78.4411
0.8998 | -78.464
0.90016 | -78.4875
0.90052 | -78.5142
0.90089 | -78.5369
0.90125 | -78.5556
0.90161 | -78.5762
0.90197 | -78.5988
0.90233 | -78.6245
0.90269 | -78.6489
0.90306 | -78.671
0.90342 | -78.6939
0.90378 | -78.7169
0.90414 | -78.7353
0.9045 | -78.7542
0.90486 | -78.7769
0.90523 | -78.8025
0.90559 | -78.828
0.90595 | -78.8473
0.90631 | -78.8644
0.90667 | -78.8851
0.90704 | -78.9088
0.9074 | -78.9344
0.90776 | -78.9598
0.90812 | -78.9761
0.90848 | -79.0028
0.90884 | -79.0305
0.90921 | -79.0573
0.90957 | -79.0833
0.90993 | -79.11
0.91029 | -79.1319
0.91065 | -79.1466
0.91102 | -79.1642
0.91138 | -79.1858
0.91174 | -79.2104
0.9121 | -79.2314
0.91246 | -79.2494
0.91282 | -79.2663
0.91319 | -79.2846
0.91355 | -79.3003
0.91391 | -79.316
0.91427 | -79.3343
0.91463 | -79.3555
0.915 | -79.3705
0.91536 | -79.3812
0.91572 | -79.3934
0.91608 | -79.4063
0.91644 | -79.4157
0.9168 | -79.4254
0.91717 | -79.4373
0.91753 | -79.4514
0.91789 | -79.4671
0.91825 | -79.4847
0.91861 | -79.4908
0.91897 | -79.5067
0.91934 | -79.521
0.9197 | -79.5304
0.92006 | -79.5331
0.92042 | -79.5369
0.92078 | -79.5377
0.92114 | -79.5402
0.92151 | -79.5441
0.92187 | -79.5534
0.92223 | -79.5617
0.92259 | -79.5569
0.92295 | -79.5533
0.92332 | -79.5598
0.92368 | -79.57
0.92404 | -79.5737
0.9244 | -79.5697
0.92476 | -79.5682
0.92512 | -79.5677
0.92549 | -79.5664
0.92585 | -79.5683
0.92621 | -79.5682
0.92657 | -79.5652
0.92693 | -79.5631
0.9273 | -79.5624
0.92766 | -79.5601
0.92802 | -79.5604
0.92838 | -79.5614
0.92874 | -79.5469
0.9291 | -79.5439
0.92947 | -79.545
0.92983 | -79.5435
0.93019 | -79.5366
0.93055 | -79.5293
0.93091 | -79.5214
0.93128 | -79.5165
0.93164 | -79.5141
0.932 | -79.5095
0.93236 | -79.5016
0.93272 | -79.4922
0.93309 | -79.4849
0.93345 | -79.4787
0.93381 | -79.4728
0.93417 | -79.4681
0.93453 | -79.4665
0.93489 | -79.4644
0.93526 | -79.459
0.93562 | -79.4535
0.93598 | -79.4499
0.93634 | -79.4463
0.9367 | -79.4432
0.93707 | -79.4412
0.93743 | -79.4388
0.93779 | -79.434
0.93815 | -79.4267
0.93851 | -79.4175
0.93887 | -79.4088
0.93924 | -79.4119
0.9396 | -79.4081
0.93996 | -79.4052
0.94032 | -79.403
0.94068 | -79.4011
0.94105 | -79.3992
0.94141 | -79.4003
0.94177 | -79.4016
0.94213 | -79.4009
0.94249 | -79.3978
0.94285 | -79.3977
0.94322 | -79.4003
0.94358 | -79.4006
0.94394 | -79.3942
0.9443 | -79.3888
0.94466 | -79.3917
0.94502 | -79.3961
0.94539 | -79.3938
0.94575 | -79.3899
0.94611 | -79.3916
0.94647 | -79.3958
0.94683 | -79.3977
0.9472 | -79.4026
0.94756 | -79.4068
0.94792 | -79.4083
0.94828 | -79.4092
0.94864 | -79.4122
0.949 | -79.4143
0.94937 | -79.4142
0.94973 | -79.4243
0.95009 | -79.4316
0.95045 | -79.4369
0.95081 | -79.4414
0.95117 | -79.4473
0.95154 | -79.4568
0.9519 | -79.4641
0.95226 | -79.4676
0.95262 | -79.4678
0.95298 | -79.4721
0.95335 | -79.4793
0.95371 | -79.486
0.95407 | -79.4898
0.95443 | -79.4957
0.95479 | -79.5014
0.95515 | -79.506
0.95552 | -79.509
0.95588 | -79.5153
0.95624 | -79.5242
0.9566 | -79.5327
0.95696 | -79.5419
0.95733 | -79.5482
0.95769 | -79.5527
0.95805 | -79.5602
0.95841 | -79.57
0.95877 | -79.5791
0.95913 | -79.5892
0.9595 | -79.5992
0.95986 | -79.6154
0.96022 | -79.6214
0.96058 | -79.6295
0.96094 | -79.6335
0.9613 | -79.64
0.96167 | -79.6464
0.96203 | -79.6529
0.96239 | -79.6601
0.96275 | -79.6705
0.96311 | -79.6818
0.96348 | -79.6889
0.96384 | -79.6936
0.9642 | -79.6981
0.96456 | -79.7054
0.96492 | -79.7122
0.96528 | -79.7134
0.96565 | -79.7174
0.96601 | -79.7178
0.96637 | -79.7071
0.96673 | -79.7069
0.96709 | -79.7103
0.96746 | -79.7122
0.96782 | -79.7123
0.96818 | -79.7153
0.96854 | -79.7199
0.9689 | -79.7205
0.96927 | -79.7168
0.96963 | -79.7211
0.96999 | -79.7268
0.97035 | -79.7285
0.97071 | -79.7263
0.97107 | -79.7253
0.97144 | -79.7236
0.9718 | -79.7266
0.97216 | -79.7321
0.97252 | -79.7328
0.97288 | -79.7314
0.97325 | -79.7302
0.97361 | -79.7306
0.97397 | -79.7318
0.97433 | -79.7315
0.97469 | -79.7299
0.97505 | -79.7284
0.97542 | -79.728
0.97578 | -79.7267
0.97614 | -79.7214
0.9765 | -79.715
0.97686 | -79.7125
0.97723 | -79.7091
0.97759 | -79.7004
0.97795 | -79.6915
0.97831 | -79.688
0.97867 | -79.6828
0.97903 | -79.6757
0.9794 | -79.6679
0.97976 | -79.6616
0.98012 | -79.6588
0.98048 | -79.6543
0.98084 | -79.6489
0.9812 | -79.6479
0.98157 | -79.6438
0.98193 | -79.6448
0.98229 | -79.6474
0.98265 | -79.6476
0.98301 | -79.6475
0.98337 | -79.6463
0.98374 | -79.6403
0.9841 | -79.6355
0.98446 | -79.6301
0.98482 | -79.6217
0.98518 | -79.6167
0.98555 | -79.6144
0.98591 | -79.6148
0.98627 | -79.6132
0.98663 | -79.6131
0.98699 | -79.6184
0.98735 | -79.6206
0.98772 | -79.6214
0.98808 | -79.6225
0.98844 | -79.6238
0.9888 | -79.6205
0.98916 | -79.6209
0.98953 | -79.6305
0.98989 | -79.6357
0.99025 | -79.6355
0.99061 | -79.6345
0.99097 | -79.6356
0.99133 | -79.6427
0.9917 | -79.6483
0.99206 | -79.6523
0.99242 | -79.6539
0.99278 | -79.6563
0.99314 | -79.6591
0.99351 | -79.6617
0.99387 | -79.6625
0.99423 | -79.6653
0.99459 | -79.6698
0.99495 | -79.6744
0.99531 | -79.6756
0.99568 | -79.681
0.99604 | -79.6907
0.9964 | -79.6965
0.99676 | -79.701
0.99712 | -79.7007
0.99748 | -79.6999
0.99785 | -79.7037
0.99821 | -79.7084
0.99857 | -79.712
0.99893 | -79.7135
0.99929 | -79.7168
0.99966 | -79.7214
1 | -79.7263
1.0004 | -79.7328
1.0007 | -79.7409
1.0011 | -79.7452
1.0015 | -79.7481
1.0018 | -79.7505
1.0022 | -79.7504
1.0026 | -79.7556
1.0029 | -79.7621
1.0033 | -79.7689
1.0036 | -79.7741
1.004 | -79.7801
1.0044 | -79.7889
1.0047 | -79.7963
1.0051 | -79.8008
1.0054 | -79.8069
1.0058 | -79.8135
1.0062 | -79.8205
1.0065 | -79.8311
1.0069 | -79.8415
1.0073 | -79.8488
1.0076 | -79.8553
1.008 | -79.8622
1.0083 | -79.8691
1.0087 | -79.8766
1.0091 | -79.8842
1.0094 | -79.8936
1.0098 | -79.9042
1.0101 | -79.911
1.0105 | -79.9137
1.0109 | -79.923
1.0112 | -79.9333
1.0116 | -79.9396
1.012 | -79.9405
1.0123 | -79.9437
1.0127 | -79.9513
1.013 | -79.957
1.0134 | -79.9599
1.0138 | -79.9632
1.0141 | -79.9695
1.0145 | -79.9768
1.0149 | -79.9805
1.0152 | -79.9834
1.0156 | -79.9902
1.0159 | -79.9936
1.0163 | -79.996
1.0167 | -80.0026
1.017 | -80.0101
1.0174 | -80.0159
1.0177 | -80.0225
1.0181 | -80.0294
1.0185 | -80.0333
1.0188 | -80.0356
1.0192 | -80.0396
1.0196 | -80.0462
1.0199 | -80.0509
1.0203 | -80.0535
1.0206 | -80.0548
1.021 | -80.055
1.0214 | -80.0569
1.0217 | -80.0643
1.0221 | -80.069
1.0225 | -80.0708
1.0228 | -80.07
1.0232 | -80.0754
1.0235 | -80.085
1.0239 | -80.0924
1.0243 | -80.0965
1.0246 | -80.1016
1.025 | -80.1095
1.0253 | -80.1141
1.0257 | -80.1122
1.0261 | -80.1105
1.0264 | -80.1106
1.0268 | -80.1146
1.0272 | -80.1195
1.0275 | -80.1203
1.0279 | -80.1205
1.0282 | -80.123
1.0286 | -80.1267
1.029 | -80.1291
1.0293 | -80.1287
1.0297 | -80.1299
1.03 | -80.1344
1.0304 | -80.1408
1.0308 | -80.1465
1.0311 | -80.1436
1.0315 | -80.1408
1.0319 | -80.1422
1.0322 | -80.1447
1.0326 | -80.1471
1.0329 | -80.1484
1.0333 | -80.1487
1.0337 | -80.1438
1.034 | -80.1452
1.0344 | -80.1494
1.0348 | -80.1521
1.0351 | -80.1548
1.0355 | -80.1622
1.0358 | -80.1677
1.0362 | -80.1754
1.0366 | -80.1928
1.0369 | -80.2016
1.0373 | -80.205
1.0376 | -80.2092
1.038 | -80.2147
1.0384 | -80.2156
1.0387 | -80.2102
1.0391 | -80.2081
1.0395 | -80.2167
1.0398 | -80.2227
1.0402 | -80.2263
1.0405 | -80.2291
1.0409 | -80.2314
1.0413 | -80.2337
1.0416 | -80.2376
1.042 | -80.2431
1.0423 | -80.2477
1.0427 | -80.2484
1.0431 | -80.2495
1.0434 | -80.2503
1.0438 | -80.2569
1.0442 | -80.2608
1.0445 | -80.2628
1.0449 | -80.2683
1.0452 | -80.2746
1.0456 | -80.2762
1.046 | -80.2808
1.0463 | -80.2881
1.0467 | -80.2961
1.0471 | -80.3019
1.0474 | -80.3114
1.0478 | -80.3234
1.0481 | -80.3311
1.0485 | -80.3403
1.0489 | -80.3503
1.0492 | -80.357
1.0496 | -80.3647
1.0499 | -80.3751
1.0503 | -80.3849
1.0507 | -80.3954
1.051 | -80.402
1.0514 | -80.4045
1.0518 | -80.4107
1.0521 | -80.4176
1.0525 | -80.4255
1.0528 | -80.4314
1.0532 | -80.4374
1.0536 | -80.4417
1.0539 | -80.4488
1.0543 | -80.457
1.0547 | -80.4632
1.055 | -80.4725
1.0554 | -80.4838
1.0557 | -80.4905
1.0561 | -80.4949
1.0565 | -80.5046
1.0568 | -80.5134
1.0572 | -80.5202
1.0575 | -80.5267
1.0579 | -80.5368
1.0583 | -80.5436
1.0586 | -80.5496
1.059 | -80.5563
1.0594 | -80.5631
1.0597 | -80.5671
1.0601 | -80.5734
1.0604 | -80.5857
1.0608 | -80.5961
1.0612 | -80.5986
1.0615 | -80.5968
1.0619 | -80.5993
1.0622 | -80.6031
1.0626 | -80.6014
1.063 | -80.6016
1.0633 | -80.6077
1.0637 | -80.6127
1.0641 | -80.6162
1.0644 | -80.6205
1.0648 | -80.6228
1.0651 | -80.6256
1.0655 | -80.6295
1.0659 | -80.6347
1.0662 | -80.6421
1.0666 | -80.6453
1.067 | -80.6461
1.0673 | -80.6537
1.0677 | -80.6624
1.068 | -80.6654
1.0684 | -80.6677
1.0688 | -80.6725
1.0691 | -80.6829
1.0695 | -80.6916
1.0698 | -80.6975
1.0702 | -80.6998
1.0706 | -80.7021
1.0709 | -80.7081
1.0713 | -80.7136
1.0717 | -80.7154
1.072 | -80.7118
1.0724 | -80.707
1.0727 | -80.7036
1.0731 | -80.6981
1.0735 | -80.6957
1.0738 | -80.6942
1.0742 | -80.6936
1.0745 | -80.6986
1.0749 | -80.698
1.0753 | -80.7033
1.0756 | -80.7024
1.076 | -80.7013
1.0764 | -80.6974
1.0767 | -80.6974
1.0771 | -80.6984
1.0774 | -80.6947
1.0778 | -80.6908
1.0782 | -80.6922
1.0785 | -80.6918
1.0789 | -80.6906
1.0793 | -80.6932
1.0796 | -80.6932
1.08 | -80.6846
1.0803 | -80.6775
1.0807 | -80.6739
1.0811 | -80.6697
1.0814 | -80.6667
1.0818 | -80.6653
1.0821 | -80.6635
1.0825 | -80.6602
1.0829 | -80.6577
1.0832 | -80.6577
1.0836 | -80.6611
1.084 | -80.6618
1.0843 | -80.6597
1.0847 | -80.6584
1.085 | -80.6622
1.0854 | -80.6596
1.0858 | -80.6628
1.0861 | -80.6647
1.0865 | -80.6687
1.0869 | -80.6731
1.0872 | -80.6777
1.0876 | -80.6803
1.0879 | -80.6818
1.0883 | -80.6772
1.0887 | -80.6718
1.089 | -80.6701
1.0894 | -80.6678
1.0897 | -80.6663
1.0901 | -80.6663
1.0905 | -80.6659
1.0908 | -80.6653
1.0912 | -80.6637
1.0916 | -80.6584
1.0919 | -80.657
1.0923 | -80.6593
1.0926 | -80.6612
1.093 | -80.662
1.0934 | -80.6585
1.0937 | -80.6571
1.0941 | -80.6582
1.0944 | -80.6595
1.0948 | -80.662
1.0952 | -80.6658
1.0955 | -80.6645
1.0959 | -80.6657
1.0963 | -80.664
1.0966 | -80.6643
1.097 | -80.6672
1.0973 | -80.6685
1.0977 | -80.6693
1.0981 | -80.6712
1.0984 | -80.6725
1.0988 | -80.6761
1.0992 | -80.6789
1.0995 | -80.682
1.0999 | -80.6812
1.1002 | -80.6794
1.1006 | -80.6842
1.101 | -80.6881
1.1013 | -80.6886
1.1017 | -80.6864
1.102 | -80.6882
1.1024 | -80.6906
1.1028 | -80.6917
1.1031 | -80.6928
1.1035 | -80.6958
1.1039 | -80.6979
1.1042 | -80.6993
1.1046 | -80.7018
1.1049 | -80.704
1.1053 | -80.7078
1.1057 | -80.7172
1.106 | -80.7262
1.1064 | -80.7294
1.1067 | -80.7334
1.1071 | -80.7351
1.1075 | -80.7381
1.1078 | -80.7432
1.1082 | -80.7478
1.1086 | -80.754
1.1089 | -80.7619
1.1093 | -80.7648
1.1096 | -80.7679
1.11 | -80.771
1.1104 | -80.7744
1.1107 | -80.7784
1.1111 | -80.7827
1.1115 | -80.789
1.1118 | -80.7933
1.1122 | -80.796
1.1125 | -80.8
1.1129 | -80.8054
1.1133 | -80.8095
1.1136 | -80.8166
1.114 | -80.8299
1.1143 | -80.8404
1.1147 | -80.8453
1.1151 | -80.8492
1.1154 | -80.8525
1.1158 | -80.8521
1.1162 | -80.8536
1.1165 | -80.8606
1.1169 | -80.8631
1.1172 | -80.8654
1.1176 | -80.8663
1.118 | -80.8689
1.1183 | -80.8753
1.1187 | -80.8818
1.1191 | -80.8888
1.1194 | -80.8962
1.1198 | -80.9002
1.1201 | -80.9044
1.1205 | -80.9105
1.1209 | -80.9149
1.1212 | -80.9171
1.1216 | -80.9253
1.1219 | -80.9285
1.1223 | -80.9327
1.1227 | -80.9359
1.123 | -80.9444
1.1234 | -80.953
1.1238 | -80.956
1.1241 | -80.9565
1.1245 | -80.9603
1.1248 | -80.9651
1.1252 | -80.9697
1.1256 | -80.9726
1.1259 | -80.9756
1.1263 | -80.9799
1.1266 | -80.9849
1.127 | -80.9915
1.1274 | -80.9986
1.1277 | -81.0071
1.1281 | -81.007
1.1285 | -81.0053
1.1288 | -81.0064
1.1292 | -81.0124
1.1295 | -81.0175
1.1299 | -81.0239
1.1303 | -81.0266
1.1306 | -81.0301
1.131 | -81.0355
1.1314 | -81.0421
1.1317 | -81.0454
1.1321 | -81.0502
1.1324 | -81.0595
1.1328 | -81.072
1.1332 | -81.0817
1.1335 | -81.0885
1.1339 | -81.0942
1.1342 | -81.0957
1.1346 | -81.0994
1.135 | -81.1046
1.1353 | -81.1071
1.1357 | -81.1052
1.1361 | -81.1106
1.1364 | -81.1193
1.1368 | -81.1263
1.1371 | -81.1323
1.1375 | -81.1422
1.1379 | -81.1465
1.1382 | -81.1533
1.1386 | -81.1654
1.1389 | -81.1741
1.1393 | -81.1838
1.1397 | -81.1919
1.14 | -81.1973
1.1404 | -81.2016
1.1408 | -81.2049
1.1411 | -81.2084
1.1415 | -81.2128
1.1418 | -81.2175
1.1422 | -81.2254
1.1426 | -81.2329
1.1429 | -81.2393
1.1433 | -81.2426
1.1437 | -81.2449
1.144 | -81.2492
1.1444 | -81.2566
1.1447 | -81.2621
1.1451 | -81.2688
1.1455 | -81.2755
1.1458 | -81.2792
1.1462 | -81.2835
1.1465 | -81.2878
1.1469 | -81.2931
1.1473 | -81.2959
1.1476 | -81.3002
1.148 | -81.3055
1.1484 | -81.3108
1.1487 | -81.3161
1.1491 | -81.323
1.1494 | -81.3309
1.1498 | -81.3427
1.1502 | -81.354
1.1505 | -81.3625
1.1509 | -81.3728
1.1513 | -81.3802
1.1516 | -81.3814
1.152 | -81.383
1.1523 | -81.3855
1.1527 | -81.3885
1.1531 | -81.3909
1.1534 | -81.3936
1.1538 | -81.3943
1.1541 | -81.3966
1.1545 | -81.401
1.1549 | -81.4016
1.1552 | -81.4032
1.1556 | -81.4085
1.156 | -81.4127
1.1563 | -81.4104
1.1567 | -81.4061
1.157 | -81.4061
1.1574 | -81.4066
1.1578 | -81.4081
1.1581 | -81.4163
1.1585 | -81.4139
1.1588 | -81.4193
1.1592 | -81.4205
1.1596 | -81.4226
1.1599 | -81.4256
1.1603 | -81.4265
1.1607 | -81.4256
1.161 | -81.4257
1.1614 | -81.4262
1.1617 | -81.4262
1.1621 | -81.4281
1.1625 | -81.4297
1.1628 | -81.4308
1.1632 | -81.4287
1.1636 | -81.4289
1.1639 | -81.4326
1.1643 | -81.4339
1.1646 | -81.4329
1.165 | -81.4315
1.1654 | -81.4325
1.1657 | -81.4311
1.1661 | -81.4291
1.1664 | -81.427
1.1668 | -81.4273
1.1672 | -81.4327
1.1675 | -81.4379
1.1679 | -81.44
1.1683 | -81.4407
1.1686 | -81.4424
1.169 | -81.4455
1.1693 | -81.4439
1.1697 | -81.4358
1.1701 | -81.432
1.1704 | -81.4321
1.1708 | -81.4355
1.1711 | -81.4399
1.1715 | -81.4387
1.1719 | -81.4371
1.1722 | -81.4346
1.1726 | -81.4357
1.173 | -81.4357
1.1733 | -81.4357
1.1737 | -81.4359
1.174 | -81.4375
1.1744 | -81.4379
1.1748 | -81.4422
1.1751 | -81.446
1.1755 | -81.4465
1.1759 | -81.4421
1.1762 | -81.4388
1.1766 | -81.4397
1.1769 | -81.4412
1.1773 | -81.4411
1.1777 | -81.4411
1.178 | -81.4421
1.1784 | -81.4422
1.1787 | -81.4443
1.1791 | -81.4456
1.1795 | -81.4516
1.1798 | -81.4551
1.1802 | -81.4626
1.1806 | -81.4624
1.1809 | -81.4622
1.1813 | -81.466
1.1816 | -81.4739
1.182 | -81.4783
1.1824 | -81.4819
1.1827 | -81.4841
1.1831 | -81.4864
1.1835 | -81.4901
1.1838 | -81.4981
1.1842 | -81.5034
1.1845 | -81.5021
1.1849 | -81.4956
1.1853 | -81.4928
1.1856 | -81.489
1.186 | -81.4854
1.1863 | -81.4862
1.1867 | -81.4858
1.1871 | -81.4857
1.1874 | -81.4903
1.1878 | -81.4942
1.1882 | -81.4967
1.1885 | -81.5002
1.1889 | -81.5062
1.1892 | -81.5134
1.1896 | -81.5132
1.19 | -81.5099
1.1903 | -81.5103
1.1907 | -81.507
1.191 | -81.5063
1.1914 | -81.5041
1.1918 | -81.5035
1.1921 | -81.5065
1.1925 | -81.5098
1.1929 | -81.5115
1.1932 | -81.5127
1.1936 | -81.5147
1.1939 | -81.5176
1.1943 | -81.5185
1.1947 | -81.5195
1.195 | -81.5236
1.1954 | -81.5282
1.1958 | -81.5299
1.1961 | -81.5304
1.1965 | -81.5354
1.1968 | -81.542
1.1972 | -81.5418
1.1976 | -81.5427
1.1979 | -81.55
1.1983 | -81.5519
1.1986 | -81.5563
1.199 | -81.5627
1.1994 | -81.5675
1.1997 | -81.569
1.2001 | -81.5709
1.2005 | -81.5742
1.2008 | -81.5735
1.2012 | -81.5734
1.2015 | -81.5742
1.2019 | -81.5764
1.2023 | -81.5781
1.2026 | -81.581
1.203 | -81.5817
1.2033 | -81.5869
1.2037 | -81.5883
1.2041 | -81.5884
1.2044 | -81.5836
1.2048 | -81.582
1.2052 | -81.582
1.2055 | -81.5847
1.2059 | -81.585
1.2062 | -81.5844
1.2066 | -81.5858
1.207 | -81.5875
1.2073 | -81.585
1.2077 | -81.5813
1.2081 | -81.5788
1.2084 | -81.577
1.2088 | -81.5704
1.2091 | -81.5647
1.2095 | -81.5642
1.2099 | -81.569
1.2102 | -81.5752
1.2106 | -81.5719
1.2109 | -81.5704
1.2113 | -81.5688
1.2117 | -81.5671
1.212 | -81.5642
1.2124 | -81.5634
1.2128 | -81.5641
1.2131 | -81.5642
1.2135 | -81.5619
1.2138 | -81.5589
1.2142 | -81.5554
1.2146 | -81.5518
1.2149 | -81.5504
1.2153 | -81.547
1.2157 | -81.54
1.216 | -81.5356
1.2164 | -81.5376
1.2167 | -81.5351
1.2171 | -81.5343
1.2175 | -81.5356
1.2178 | -81.5418
1.2182 | -81.5436
1.2185 | -81.5409
1.2189 | -81.5363
1.2193 | -81.5324
1.2196 | -81.5285
1.22 | -81.524
1.2204 | -81.5222
1.2207 | -81.5228
1.2211 | -81.5193
1.2214 | -81.518
1.2218 | -81.5162
1.2222 | -81.5127
1.2225 | -81.5103
1.2229 | -81.5093
1.2232 | -81.5076
1.2236 | -81.5042
1.224 | -81.4993
1.2243 | -81.4942
1.2247 | -81.4908
1.2251 | -81.4882
1.2254 | -81.4855
1.2258 | -81.483
1.2261 | -81.4831
1.2265 | -81.4846
1.2269 | -81.4856
1.2272 | -81.4842
1.2276 | -81.4818
1.228 | -81.4803
1.2283 | -81.4801
1.2287 | -81.48
1.229 | -81.4768
1.2294 | -81.4724
1.2298 | -81.471
1.2301 | -81.4701
1.2305 | -81.466
1.2308 | -81.4647
1.2312 | -81.4638
1.2316 | -81.4673
1.2319 | -81.4665
1.2323 | -81.4641
1.2327 | -81.463
1.233 | -81.4634
1.2334 | -81.4643
1.2337 | -81.4644
1.2341 | -81.4639
1.2345 | -81.4637
1.2348 | -81.4658
1.2352 | -81.4714
1.2356 | -81.4714
1.2359 | -81.4699
1.2363 | -81.4715
1.2366 | -81.4735
1.237 | -81.4742
1.2374 | -81.4779
1.2377 | -81.4792
1.2381 | -81.4776
1.2384 | -81.4816
1.2388 | -81.4848
1.2392 | -81.4862
1.2395 | -81.4883
1.2399 | -81.4952
1.2403 | -81.5034
1.2406 | -81.5075
1.241 | -81.5122
1.2413 | -81.5155
1.2417 | -81.5164
1.2421 | -81.5227
1.2424 | -81.522
1.2428 | -81.5255
1.2431 | -81.5268
1.2435 | -81.5289
1.2439 | -81.5337
1.2442 | -81.5365
1.2446 | -81.5402
1.245 | -81.5429
1.2453 | -81.5442
1.2457 | -81.5474
1.246 | -81.5525
1.2464 | -81.5571
1.2468 | -81.5616
1.2471 | -81.5672
1.2475 | -81.5751
1.2479 | -81.5812
1.2482 | -81.5875
1.2486 | -81.5927
1.2489 | -81.599
1.2493 | -81.6044
1.2497 | -81.6089
1.25 | -81.6162
1.2504 | -81.6189
1.2507 | -81.6227
1.2511 | -81.6286
1.2515 | -81.6329
1.2518 | -81.6339
1.2522 | -81.6361
1.2526 | -81.6441
1.2529 | -81.6519
1.2533 | -81.6561
1.2536 | -81.6577
1.254 | -81.6605
1.2544 | -81.667
1.2547 | -81.681
1.2551 | -81.69
1.2554 | -81.6971
1.2558 | -81.7036
1.2562 | -81.7085
1.2565 | -81.7155
1.2569 | -81.7216
1.2573 | -81.7256
1.2576 | -81.7291
1.258 | -81.7322
1.2583 | -81.7373
1.2587 | -81.7437
1.2591 | -81.7467
1.2594 | -81.7482
1.2598 | -81.7509
1.2602 | -81.756
1.2605 | -81.762
1.2609 | -81.7725
1.2612 | -81.7773
1.2616 | -81.7805
1.262 | -81.7851
1.2623 | -81.7896
1.2627 | -81.7945
1.263 | -81.7991
1.2634 | -81.8024
1.2638 | -81.8043
1.2641 | -81.8074
1.2645 | -81.8093
1.2649 | -81.808
1.2652 | -81.8094
1.2656 | -81.8156
1.2659 | -81.8205
1.2663 | -81.8251
1.2667 | -81.8294
1.267 | -81.8323
1.2674 | -81.8302
1.2678 | -81.8307
1.2681 | -81.8329
1.2685 | -81.8351
1.2688 | -81.8366
1.2692 | -81.8382
1.2696 | -81.839
1.2699 | -81.8386
1.2703 | -81.8426
1.2706 | -81.8513
1.271 | -81.8541
1.2714 | -81.8584
1.2717 | -81.8622
1.2721 | -81.8653
1.2725 | -81.8688
1.2728 | -81.8706
1.2732 | -81.869
1.2735 | -81.8676
1.2739 | -81.8657
1.2743 | -81.8645
1.2746 | -81.864
1.275 | -81.8627
1.2753 | -81.861
1.2757 | -81.862
1.2761 | -81.8647
1.2764 | -81.8603
1.2768 | -81.8563
1.2772 | -81.8557
1.2775 | -81.853
1.2779 | -81.8525
1.2782 | -81.8519
1.2786 | -81.8479
1.279 | -81.8429
1.2793 | -81.8387
1.2797 | -81.836
1.2801 | -81.8329
1.2804 | -81.8291
1.2808 | -81.827
1.2811 | -81.8246
1.2815 | -81.8191
1.2819 | -81.8142
1.2822 | -81.8115
1.2826 | -81.811
1.2829 | -81.8109
1.2833 | -81.8074
1.2837 | -81.8075
1.284 | -81.8069
1.2844 | -81.8061
1.2848 | -81.8015
1.2851 | -81.8002
1.2855 | -81.7948
1.2858 | -81.7815
1.2862 | -81.7761
1.2866 | -81.7755
1.2869 | -81.7732
1.2873 | -81.7711
1.2876 | -81.7696
1.288 | -81.7656
1.2884 | -81.7607
1.2887 | -81.7582
1.2891 | -81.7565
1.2895 | -81.7551
1.2898 | -81.7533
1.2902 | -81.7499
1.2905 | -81.7474
1.2909 | -81.7462
1.2913 | -81.7473
1.2916 | -81.7539
1.292 | -81.755
1.2924 | -81.7531
1.2927 | -81.7529
1.2931 | -81.7542
1.2934 | -81.755
1.2938 | -81.7548
1.2942 | -81.7586
1.2945 | -81.7584
1.2949 | -81.7612
1.2952 | -81.7648
1.2956 | -81.7667
1.296 | -81.768
1.2963 | -81.7741
1.2967 | -81.7787
1.2971 | -81.78
1.2974 | -81.7813
1.2978 | -81.7818
1.2981 | -81.7831
1.2985 | -81.7874
1.2989 | -81.7912
1.2992 | -81.7921
1.2996 | -81.7937
1.3 | -81.7954
1.3003 | -81.7962
1.3007 | -81.7979
1.301 | -81.7996
1.3014 | -81.8016
1.3018 | -81.8057
1.3021 | -81.8103
1.3025 | -81.8142
1.3028 | -81.8127
1.3032 | -81.8119
1.3036 | -81.8189
1.3039 | -81.8261
1.3043 | -81.8315
1.3047 | -81.841
1.305 | -81.8455
1.3054 | -81.8487
1.3057 | -81.8545
1.3061 | -81.8598
1.3065 | -81.8647
1.3068 | -81.8706
1.3072 | -81.8776
1.3075 | -81.8833
1.3079 | -81.8885
1.3083 | -81.8951
1.3086 | -81.9012
1.309 | -81.9057
1.3094 | -81.9101
1.3097 | -81.9152
1.3101 | -81.9195
1.3104 | -81.9252
1.3108 | -81.9312
1.3112 | -81.9388
1.3115 | -81.9468
1.3119 | -81.9525
1.3123 | -81.9601
1.3126 | -81.9696
1.313 | -81.9775
1.3133 | -81.9845
1.3137 | -81.9912
1.3141 | -81.9969
1.3144 | -81.9987
1.3148 | -82.0041
1.3151 | -82.0086
1.3155 | -82.0143
1.3159 | -82.0207
1.3162 | -82.0293
1.3166 | -82.0372
1.317 | -82.043
1.3173 | -82.0494
1.3177 | -82.0573
1.318 | -82.0676
1.3184 | -82.0753
1.3188 | -82.0797
1.3191 | -82.0876
1.3195 | -82.0973
1.3198 | -82.1051
1.3202 | -82.1107
1.3206 | -82.1163
1.3209 | -82.1213
1.3213 | -82.1252
1.3217 | -82.1337
1.322 | -82.1404
1.3224 | -82.1468
1.3227 | -82.1531
1.3231 | -82.1614
1.3235 | -82.1684
1.3238 | -82.1734
1.3242 | -82.1802
1.3246 | -82.1885
1.3249 | -82.2007
1.3253 | -82.2106
1.3256 | -82.2164
1.326 | -82.2218
1.3264 | -82.2249
1.3267 | -82.228
1.3271 | -82.2333
1.3274 | -82.2413
1.3278 | -82.2515
1.3282 | -82.2561
1.3285 | -82.2606
1.3289 | -82.2674
1.3293 | -82.2746
1.3296 | -82.2802
1.33 | -82.286
1.3303 | -82.2911
1.3307 | -82.2935
1.3311 | -82.2984
1.3314 | -82.3063
1.3318 | -82.3127
1.3322 | -82.32
1.3325 | -82.3307
1.3329 | -82.3381
1.3332 | -82.3465
1.3336 | -82.3512
1.334 | -82.3524
1.3343 | -82.358
1.3347 | -82.362
1.335 | -82.365
1.3354 | -82.3668
1.3358 | -82.3686
1.3361 | -82.3707
1.3365 | -82.3737
1.3369 | -82.3753
1.3372 | -82.3734
1.3376 | -82.3702
1.3379 | -82.3705
1.3383 | -82.3748
1.3387 | -82.3776
1.339 | -82.3809
1.3394 | -82.382
1.3397 | -82.3852
1.3401 | -82.3881
1.3405 | -82.39
1.3408 | -82.3958
1.3412 | -82.4015
1.3416 | -82.404
1.3419 | -82.4057
1.3423 | -82.409
1.3426 | -82.4096
1.343 | -82.41
1.3434 | -82.4113
1.3437 | -82.4124
1.3441 | -82.4156
1.3445 | -82.4202
1.3448 | -82.4218
1.3452 | -82.4169
1.3455 | -82.4089
1.3459 | -82.4069
1.3463 | -82.4039
1.3466 | -82.4031
1.347 | -82.4039
1.3473 | -82.4037
1.3477 | -82.4016
1.3481 | -82.3974
1.3484 | -82.4004
1.3488 | -82.4048
1.3492 | -82.4073
1.3495 | -82.4042
1.3499 | -82.4035
1.3502 | -82.4059
1.3506 | -82.4105
1.351 | -82.413
1.3513 | -82.4109
1.3517 | -82.4083
1.3521 | -82.4046
1.3524 | -82.4037
1.3528 | -82.4021
1.3531 | -82.3989
1.3535 | -82.3959
1.3539 | -82.393
1.3542 | -82.388
1.3546 | -82.3837
1.3549 | -82.3799
1.3553 | -82.3746
1.3557 | -82.3674
1.356 | -82.3638
1.3564 | -82.3647
1.3568 | -82.3617
1.3571 | -82.3606
1.3575 | -82.3567
1.3578 | -82.3519
1.3582 | -82.345
1.3586 | -82.3405
1.3589 | -82.3368
1.3593 | -82.3364
1.3596 | -82.3345
1.36 | -82.3298
1.3604 | -82.3256
1.3607 | -82.3219
1.3611 | -82.3179
1.3615 | -82.3153
1.3618 | -82.3134
1.3622 | -82.3104
1.3625 | -82.3081
1.3629 | -82.306
1.3633 | -82.3024
1.3636 | -82.2995
1.364 | -82.2955
1.3644 | -82.289
1.3647 | -82.2771
1.3651 | -82.2687
1.3654 | -82.2636
1.3658 | -82.2579
1.3662 | -82.251
1.3665 | -82.2455
1.3669 | -82.2378
1.3672 | -82.2314
1.3676 | -82.2241
1.368 | -82.2176
1.3683 | -82.2144
1.3687 | -82.2057
1.3691 | -82.1997
1.3694 | -82.1957
1.3698 | -82.1939
1.3701 | -82.1906
1.3705 | -82.1862
1.3709 | -82.1802
1.3712 | -82.1724
1.3716 | -82.169
1.372 | -82.1659
1.3723 | -82.1607
1.3727 | -82.1546
1.373 | -82.1496
1.3734 | -82.1445
1.3738 | -82.1389
1.3741 | -82.1352
1.3745 | -82.1289
1.3748 | -82.1228
1.3752 | -82.118
1.3756 | -82.1123
1.3759 | -82.1074
1.3763 | -82.103
1.3767 | -82.0983
1.377 | -82.0932
1.3774 | -82.0863
1.3777 | -82.0785
1.3781 | -82.0761
1.3785 | -82.0737
1.3788 | -82.0726
1.3792 | -82.0764
1.3795 | -82.0731
1.3799 | -82.0658
1.3803 | -82.0586
1.3806 | -82.0531
1.381 | -82.0483
1.3814 | -82.0483
1.3817 | -82.0496
1.3821 | -82.0455
1.3824 | -82.039
1.3828 | -82.0338
1.3832 | -82.0272
1.3835 | -82.0191
1.3839 | -82.0098
1.3843 | -82.0056
1.3846 | -82.0029
1.385 | -81.9997
1.3853 | -81.9966
1.3857 | -81.9964
1.3861 | -81.9935
1.3864 | -81.9906
1.3868 | -81.9888
1.3871 | -81.9878
1.3875 | -81.9872
1.3879 | -81.9869
1.3882 | -81.9896
1.3886 | -81.9897
1.389 | -81.9908
1.3893 | -81.9904
1.3897 | -81.9907
1.39 | -81.9913
1.3904 | -81.9891
1.3908 | -81.9879
1.3911 | -81.9854
1.3915 | -81.9859
1.3918 | -81.9832
1.3922 | -81.9788
1.3926 | -81.9775
1.3929 | -81.9781
1.3933 | -81.9749
1.3937 | -81.9704
1.394 | -81.9718
1.3944 | -81.9749
1.3947 | -81.9753
1.3951 | -81.9765
1.3955 | -81.9792
1.3958 | -81.9809
1.3962 | -81.9815
1.3966 | -81.9833
1.3969 | -81.9879
1.3973 | -81.9949
1.3976 | -81.9966
1.398 | -81.9983
1.3984 | -82.0002
1.3987 | -81.9977
1.3991 | -81.9986
1.3994 | -82.0016
1.3998 | -82.0022
1.4002 | -82.0003
1.4005 | -81.9972
1.4009 | -81.9876
1.4013 | -81.9797
1.4016 | -81.981
1.402 | -81.9833
1.4023 | -81.9858
1.4027 | -81.9873
1.4031 | -81.9877
1.4034 | -81.985
1.4038 | -81.9852
1.4042 | -81.9893
1.4045 | -81.9938
1.4049 | -81.9982
1.4052 | -82.0006
1.4056 | -82.003
1.406 | -82.0081
1.4063 | -82.0108
1.4067 | -82.0107
1.407 | -82.0124
1.4074 | -82.016
1.4078 | -82.0195
1.4081 | -82.02
1.4085 | -82.0201
1.4089 | -82.0218
1.4092 | -82.023
1.4096 | -82.0272
1.4099 | -82.0294
1.4103 | -82.0305
1.4107 | -82.0353
1.411 | -82.0344
1.4114 | -82.0325
1.4117 | -82.0322
1.4121 | -82.0318
1.4125 | -82.0317
1.4128 | -82.0277
1.4132 | -82.0249
1.4136 | -82.0251
1.4139 | -82.0263
1.4143 | -82.0261
1.4146 | -82.0262
1.415 | -82.0262
1.4154 | -82.0251
1.4157 | -82.0224
1.4161 | -82.0212
1.4165 | -82.0215
1.4168 | -82.0219
1.4172 | -82.0212
1.4175 | -82.0203
1.4179 | -82.019
1.4183 | -82.0185
1.4186 | -82.0182
1.419 | -82.0187
1.4193 | -82.0186
1.4197 | -82.019
1.4201 | -82.0186
1.4204 | -82.0182
1.4208 | -82.0173
1.4212 | -82.0166
1.4215 | -82.016
1.4219 | -82.014
1.4222 | -82.0086
1.4226 | -82.0032
1.423 | -81.9974
1.4233 | -81.9937
1.4237 | -81.991
1.424 | -81.9877
1.4244 | -81.9861
1.4248 | -81.9841
1.4251 | -81.9796
1.4255 | -81.975
1.4259 | -81.9728
1.4262 | -81.9712
1.4266 | -81.9693
1.4269 | -81.967
1.4273 | -81.9644
1.4277 | -81.9616
1.428 | -81.9578
1.4284 | -81.9526
1.4288 | -81.9494
1.4291 | -81.947
1.4295 | -81.9442
1.4298 | -81.9376
1.4302 | -81.9372
1.4306 | -81.9353
1.4309 | -81.9309
1.4313 | -81.928
1.4316 | -81.9281
1.432 | -81.9289
1.4324 | -81.9247
1.4327 | -81.9204
1.4331 | -81.9183
1.4335 | -81.9131
1.4338 | -81.9084
1.4342 | -81.9087
1.4345 | -81.9109
1.4349 | -81.9107
1.4353 | -81.9078
1.4356 | -81.904
1.436 | -81.898
1.4364 | -81.8915
1.4367 | -81.8851
1.4371 | -81.8817
1.4374 | -81.8795
1.4378 | -81.8762
1.4382 | -81.872
1.4385 | -81.8678
1.4389 | -81.8663
1.4392 | -81.8637
1.4396 | -81.8604
1.44 | -81.8554
1.4403 | -81.856
1.4407 | -81.8527
1.4411 | -81.8489
1.4414 | -81.8462
1.4418 | -81.8433
1.4421 | -81.8409
1.4425 | -81.8451
1.4429 | -81.8482
1.4432 | -81.8452
1.4436 | -81.8426
1.4439 | -81.8394
1.4443 | -81.8368
1.4447 | -81.8344
1.445 | -81.8324
1.4454 | -81.8286
1.4458 | -81.8262
1.4461 | -81.8244
1.4465 | -81.8216
1.4468 | -81.8205
1.4472 | -81.8214
1.4476 | -81.8192
1.4479 | -81.8161
1.4483 | -81.8143
1.4487 | -81.8117
1.449 | -81.8069
1.4494 | -81.8047
1.4497 | -81.8048
1.4501 | -81.804
1.4505 | -81.8033
1.4508 | -81.8059
1.4512 | -81.8049
1.4515 | -81.804
1.4519 | -81.8039
1.4523 | -81.8057
1.4526 | -81.8065
1.453 | -81.8059
1.4534 | -81.8069
1.4537 | -81.8108
1.4541 | -81.8128
1.4544 | -81.8136
1.4548 | -81.814
1.4552 | -81.8143
1.4555 | -81.8143
1.4559 | -81.814
1.4562 | -81.8153
1.4566 | -81.8166
1.457 | -81.8161
1.4573 | -81.8182
1.4577 | -81.8206
1.4581 | -81.8214
1.4584 | -81.8232
1.4588 | -81.8262
1.4591 | -81.8275
1.4595 | -81.8267
1.4599 | -81.8276
1.4602 | -81.8294
1.4606 | -81.8311
1.461 | -81.832
1.4613 | -81.8386
1.4617 | -81.8413
1.462 | -81.8421
1.4624 | -81.8442
1.4628 | -81.8462
1.4631 | -81.8494
1.4635 | -81.8564
1.4638 | -81.8611
1.4642 | -81.8623
1.4646 | -81.8661
1.4649 | -81.8713
1.4653 | -81.8747
1.4657 | -81.8803
1.466 | -81.8856
1.4664 | -81.8896
1.4667 | -81.8953
1.4671 | -81.9004
1.4675 | -81.9057
1.4678 | -81.9157
1.4682 | -81.9233
1.4686 | -81.9271
1.4689 | -81.9301
1.4693 | -81.9351
1.4696 | -81.9394
1.47 | -81.9429
1.4704 | -81.9477
1.4707 | -81.952
1.4711 | -81.9566
1.4714 | -81.9608
1.4718 | -81.9682
1.4722 | -81.9734
1.4725 | -81.9794
1.4729 | -81.9852
1.4733 | -81.9919
1.4736 | -81.998
1.474 | -82.0038
1.4743 | -82.0099
1.4747 | -82.0153
1.4751 | -82.0197
1.4754 | -82.0265
1.4758 | -82.0334
1.4761 | -82.0405
1.4765 | -82.0482
1.4769 | -82.0569
1.4772 | -82.068
1.4776 | -82.0778
1.478 | -82.0829
1.4783 | -82.0904
1.4787 | -82.0996
1.479 | -82.1081
1.4794 | -82.1128
1.4798 | -82.1221
1.4801 | -82.133
1.4805 | -82.1424
1.4809 | -82.1487
1.4812 | -82.1553
1.4816 | -82.1634
1.4819 | -82.1742
1.4823 | -82.1869
1.4827 | -82.1982
1.483 | -82.2078
1.4834 | -82.2208
1.4837 | -82.2374
1.4841 | -82.252
1.4845 | -82.2618
1.4848 | -82.2696
1.4852 | -82.2787
1.4856 | -82.2896
1.4859 | -82.2996
1.4863 | -82.3102
1.4866 | -82.3212
1.487 | -82.3315
1.4874 | -82.3412
1.4877 | -82.3518
1.4881 | -82.3622
1.4885 | -82.3735
1.4888 | -82.3831
1.4892 | -82.3943
1.4895 | -82.4027
1.4899 | -82.413
1.4903 | -82.4235
1.4906 | -82.4337
1.491 | -82.4461
1.4913 | -82.4571
1.4917 | -82.4712
1.4921 | -82.4818
1.4924 | -82.4876
1.4928 | -82.5011
1.4932 | -82.5161
1.4935 | -82.5286
1.4939 | -82.5415
1.4942 | -82.5589
1.4946 | -82.5756
1.495 | -82.5817
1.4953 | -82.586
1.4957 | -82.5925
1.496 | -82.5998
1.4964 | -82.6092
1.4968 | -82.6203
1.4971 | -82.6306
1.4975 | -82.6419
1.4979 | -82.6533
1.4982 | -82.6638
1.4986 | -82.6757
1.4989 | -82.6869
1.4993 | -82.6977
1.4997 | -82.7133
1.5 | -82.7305
1.5004 | -82.7432
1.5008 | -82.7533
1.5011 | -82.7637
1.5015 | -82.7778
1.5018 | -82.794
1.5022 | -82.8174
1.5026 | -82.837
1.5029 | -82.845
1.5033 | -82.8531
1.5036 | -82.8587
1.504 | -82.8586
1.5044 | -82.8603
1.5047 | -82.87
1.5051 | -82.8842
1.5055 | -82.8946
1.5058 | -82.9015
1.5062 | -82.9103
1.5065 | -82.9183
1.5069 | -82.9257
1.5073 | -82.9298
1.5076 | -82.9362
1.508 | -82.9439
1.5083 | -82.9536
1.5087 | -82.9587
1.5091 | -82.9635
1.5094 | -82.9702
1.5098 | -82.9791
1.5102 | -82.9861
1.5105 | -82.9893
1.5109 | -82.9939
1.5112 | -82.9994
1.5116 | -83.004
1.512 | -83.0091
1.5123 | -83.0142
1.5127 | -83.0193
1.5131 | -83.0273
1.5134 | -83.0237
1.5138 | -83.0304
1.5141 | -83.0325
1.5145 | -83.0353
1.5149 | -83.0399
1.5152 | -83.049
1.5156 | -83.0539
1.5159 | -83.054
1.5163 | -83.0557
1.5167 | -83.0596
1.517 | -83.0627
1.5174 | -83.0647
1.5178 | -83.0653
1.5181 | -83.0651
1.5185 | -83.0669
1.5188 | -83.0672
1.5192 | -83.0658
1.5196 | -83.0637
1.5199 | -83.0644
1.5203 | -83.0649
1.5207 | -83.0645
1.521 | -83.0623
1.5214 | -83.0599
1.5217 | -83.058
1.5221 | -83.0563
1.5225 | -83.0558
1.5228 | -83.0538
1.5232 | -83.0516
1.5235 | -83.049
1.5239 | -83.0438
1.5243 | -83.0414
1.5246 | -83.0379
1.525 | -83.0371
1.5254 | -83.0361
1.5257 | -83.0369
1.5261 | -83.0403
1.5264 | -83.0386
1.5268 | -83.0353
1.5272 | -83.0304
1.5275 | -83.0216
1.5279 | -83.013
1.5282 | -83.0064
1.5286 | -83.004
1.529 | -82.9971
1.5293 | -82.9875
1.5297 | -82.9796
1.5301 | -82.9748
1.5304 | -82.97
1.5308 | -82.9628
1.5311 | -82.9553
1.5315 | -82.949
1.5319 | -82.9414
1.5322 | -82.9335
1.5326 | -82.9264
1.533 | -82.9209
1.5333 | -82.9165
1.5337 | -82.9094
1.534 | -82.9037
1.5344 | -82.8943
1.5348 | -82.8887
1.5351 | -82.8806
1.5355 | -82.8694
1.5358 | -82.864
1.5362 | -82.8597
1.5366 | -82.8463
1.5369 | -82.8337
1.5373 | -82.8261
1.5377 | -82.8209
1.538 | -82.8128
1.5384 | -82.8025
1.5387 | -82.7904
1.5391 | -82.7715
1.5395 | -82.7587
1.5398 | -82.752
1.5402 | -82.7487
1.5405 | -82.7433
1.5409 | -82.7354
1.5413 | -82.7273
1.5416 | -82.7197
1.542 | -82.7106
1.5424 | -82.698
1.5427 | -82.6843
1.5431 | -82.6735
1.5434 | -82.6648
1.5438 | -82.6562
1.5442 | -82.6455
1.5445 | -82.6343
1.5449 | -82.6182
1.5453 | -82.6063
1.5456 | -82.5963
1.546 | -82.5836
1.5463 | -82.5734
1.5467 | -82.5655
1.5471 | -82.5517
1.5474 | -82.5376
1.5478 | -82.5279
1.5481 | -82.5199
1.5485 | -82.511
1.5489 | -82.5012
1.5492 | -82.4918
1.5496 | -82.4831
1.55 | -82.4758
1.5503 | -82.471
1.5507 | -82.4629
1.551 | -82.4543
1.5514 | -82.444
1.5518 | -82.4329
1.5521 | -82.4233
1.5525 | -82.4157
1.5529 | -82.4073
1.5532 | -82.3996
1.5536 | -82.3939
1.5539 | -82.3874
1.5543 | -82.3808
1.5547 | -82.3753
1.555 | -82.3661
1.5554 | -82.3578
1.5557 | -82.3493
1.5561 | -82.3396
1.5565 | -82.3332
1.5568 | -82.3255
1.5572 | -82.3174
1.5576 | -82.3099
1.5579 | -82.3032
1.5583 | -82.2979
1.5586 | -82.2915
1.559 | -82.2821
1.5594 | -82.2738
1.5597 | -82.2673
1.5601 | -82.261
1.5604 | -82.2533
1.5608 | -82.2469
1.5612 | -82.2407
1.5615 | -82.2339
1.5619 | -82.2303
1.5623 | -82.2263
1.5626 | -82.218
1.563 | -82.2126
1.5633 | -82.2078
1.5637 | -82.2026
1.5641 | -82.1954
1.5644 | -82.1873
1.5648 | -82.1842
1.5652 | -82.1801
1.5655 | -82.1743
1.5659 | -82.1678
1.5662 | -82.1618
1.5666 | -82.157
1.567 | -82.153
1.5673 | -82.1487
1.5677 | -82.1448
1.568 | -82.1424
1.5684 | -82.1347
1.5688 | -82.1262
1.5691 | -82.1187
1.5695 | -82.1119
1.5699 | -82.1049
1.5702 | -82.1018
1.5706 | -82.0974
1.5709 | -82.0928
1.5713 | -82.0896
1.5717 | -82.0842
1.572 | -82.0798
1.5724 | -82.0765
1.5727 | -82.0736
1.5731 | -82.0701
1.5735 | -82.0634
1.5738 | -82.057
1.5742 | -82.0513
1.5746 | -82.0468
1.5749 | -82.0384
1.5753 | -82.0333
1.5756 | -82.0302
1.576 | -82.0309
1.5764 | -82.0288
1.5767 | -82.027
1.5771 | -82.0253
1.5775 | -82.0209
1.5778 | -82.0176
1.5782 | -82.0167
1.5785 | -82.0149
1.5789 | -82.0133
1.5793 | -82.0106
1.5796 | -82.0058
1.58 | -82.0031
1.5803 | -82.0022
1.5807 | -82.0019
1.5811 | -81.9971
1.5814 | -81.9902
1.5818 | -81.9884
1.5822 | -81.9855
1.5825 | -81.9836
1.5829 | -81.9819
1.5832 | -81.9803
1.5836 | -81.9785
1.584 | -81.9773
1.5843 | -81.9764
1.5847 | -81.9772
1.5851 | -81.9757
1.5854 | -81.972
1.5858 | -81.9711
1.5861 | -81.9743
1.5865 | -81.9713
1.5869 | -81.9679
1.5872 | -81.9619
1.5876 | -81.9549
1.5879 | -81.9525
1.5883 | -81.9528
1.5887 | -81.9503
1.589 | -81.9482
1.5894 | -81.9444
1.5898 | -81.9411
1.5901 | -81.9378
1.5905 | -81.9342
1.5908 | -81.9287
1.5912 | -81.9251
1.5916 | -81.9235
1.5919 | -81.9224
1.5923 | -81.9217
1.5926 | -81.9178
1.593 | -81.9139
1.5934 | -81.9127
1.5937 | -81.9125
1.5941 | -81.9115
1.5945 | -81.9097
1.5948 | -81.9076
1.5952 | -81.9067
1.5955 | -81.9062
1.5959 | -81.9071
1.5963 | -81.9072
1.5966 | -81.9053
1.597 | -81.9073
1.5974 | -81.909
1.5977 | -81.9088
1.5981 | -81.9069
1.5984 | -81.9079
1.5988 | -81.9096
1.5992 | -81.9094
1.5995 | -81.9099
1.5999 | -81.9098
1.6002 | -81.912
1.6006 | -81.914
1.601 | -81.9132
1.6013 | -81.9135
1.6017 | -81.914
1.6021 | -81.9141
1.6024 | -81.9137
1.6028 | -81.9121
1.6031 | -81.9142
1.6035 | -81.9167
1.6039 | -81.9165
1.6042 | -81.9163
1.6046 | -81.9166
1.6049 | -81.9147
1.6053 | -81.913
1.6057 | -81.911
1.606 | -81.9115
1.6064 | -81.9113
1.6068 | -81.9093
1.6071 | -81.9071
1.6075 | -81.9071
1.6078 | -81.9059
1.6082 | -81.9044
1.6086 | -81.9035
1.6089 | -81.9013
1.6093 | -81.9018
1.6097 | -81.9037
1.61 | -81.9018
1.6104 | -81.8956
1.6107 | -81.8914
1.6111 | -81.8867
1.6115 | -81.8809
1.6118 | -81.8781
1.6122 | -81.8781
1.6125 | -81.8801
1.6129 | -81.8866
1.6133 | -81.8927
1.6136 | -81.8927
1.614 | -81.8887
1.6144 | -81.8873
1.6147 | -81.8874
1.6151 | -81.8861
1.6154 | -81.8835
1.6158 | -81.8828
1.6162 | -81.8844
1.6165 | -81.8841
1.6169 | -81.8828
1.6173 | -81.8809
1.6176 | -81.8817
1.618 | -81.8839
1.6183 | -81.8852
1.6187 | -81.8844
1.6191 | -81.8841
1.6194 | -81.8834
1.6198 | -81.8847
1.6201 | -81.8832
1.6205 | -81.8803
1.6209 | -81.8788
1.6212 | -81.8756
1.6216 | -81.8752
1.622 | -81.8757
1.6223 | -81.8741
1.6227 | -81.8726
1.623 | -81.8757
1.6234 | -81.8764
1.6238 | -81.8775
1.6241 | -81.8789
1.6245 | -81.8773
1.6248 | -81.8749
1.6252 | -81.8733
1.6256 | -81.8741
1.6259 | -81.8756
1.6263 | -81.8769
1.6267 | -81.8761
1.627 | -81.8755
1.6274 | -81.876
1.6277 | -81.8764
1.6281 | -81.874
1.6285 | -81.8739
1.6288 | -81.8743
1.6292 | -81.8757
1.6296 | -81.876
1.6299 | -81.8742
1.6303 | -81.8734
1.6306 | -81.8731
1.631 | -81.8744
1.6314 | -81.8751
1.6317 | -81.8745
1.6321 | -81.8728
1.6324 | -81.8709
1.6328 | -81.8674
1.6332 | -81.86
1.6335 | -81.8586
1.6339 | -81.8587
1.6343 | -81.8577
1.6346 | -81.8586
1.635 | -81.8584
1.6353 | -81.8584
1.6357 | -81.8595
1.6361 | -81.8594
1.6364 | -81.8586
1.6368 | -81.8584
1.6372 | -81.8588
1.6375 | -81.8602
1.6379 | -81.8633
1.6382 | -81.8673
1.6386 | -81.8706
1.639 | -81.8746
1.6393 | -81.8753
1.6397 | -81.8763
1.64 | -81.8792
1.6404 | -81.8817
1.6408 | -81.884
1.6411 | -81.8826
1.6415 | -81.8845
1.6419 | -81.8867
1.6422 | -81.8895
1.6426 | -81.8915
1.6429 | -81.8932
1.6433 | -81.8949
1.6437 | -81.8973
1.644 | -81.8987
1.6444 | -81.8994
1.6447 | -81.9014
1.6451 | -81.904
1.6455 | -81.9057
1.6458 | -81.9034
1.6462 | -81.9014
1.6466 | -81.9064
1.6469 | -81.9076
1.6473 | -81.9035
1.6476 | -81.9063
1.648 | -81.9093
1.6484 | -81.9108
1.6487 | -81.9107
1.6491 | -81.9091
1.6495 | -81.9098
1.6498 | -81.9111
1.6502 | -81.9143
1.6505 | -81.9228
1.6509 | -81.9256
1.6513 | -81.9248
1.6516 | -81.9257
1.652 | -81.9206
1.6523 | -81.9202
1.6527 | -81.9216
1.6531 | -81.9226
1.6534 | -81.9232
1.6538 | -81.9274
1.6542 | -81.9312
1.6545 | -81.9334
1.6549 | -81.9361
1.6552 | -81.9385
1.6556 | -81.9389
1.656 | -81.9404
1.6563 | -81.9439
1.6567 | -81.9496
1.657 | -81.9519
1.6574 | -81.9513
1.6578 | -81.9539
1.6581 | -81.9566
1.6585 | -81.9575
1.6589 | -81.9618
1.6592 | -81.9657
1.6596 | -81.9562
1.6599 | -81.958
1.6603 | -81.9572
1.6607 | -81.9505
1.661 | -81.9483
1.6614 | -81.9468
1.6618 | -81.9442
1.6621 | -81.944
1.6625 | -81.9423
1.6628 | -81.9405
1.6632 | -81.9403
1.6636 | -81.9407
1.6639 | -81.9424
1.6643 | -81.9436
1.6646 | -81.9445
1.665 | -81.9454
1.6654 | -81.9451
1.6657 | -81.9436
1.6661 | -81.941
1.6665 | -81.9415
1.6668 | -81.9379
1.6672 | -81.9345
1.6675 | -81.934
1.6679 | -81.9345
1.6683 | -81.9356
1.6686 | -81.9379
1.669 | -81.9375
1.6693 | -81.9369
1.6697 | -81.9347
1.6701 | -81.9278
1.6704 | -81.9261
1.6708 | -81.9254
1.6712 | -81.9237
1.6715 | -81.9226
1.6719 | -81.9227
1.6722 | -81.9167
1.6726 | -81.9061
1.673 | -81.9035
1.6733 | -81.9032
1.6737 | -81.9005
1.6741 | -81.9002
1.6744 | -81.8984
1.6748 | -81.8965
1.6751 | -81.8946
1.6755 | -81.8943
1.6759 | -81.8885
1.6762 | -81.8841
1.6766 | -81.8825
1.6769 | -81.8818
1.6773 | -81.8823
1.6777 | -81.8795
1.678 | -81.8761
1.6784 | -81.8736
1.6788 | -81.8711
1.6791 | -81.869
1.6795 | -81.8666
1.6798 | -81.863
1.6802 | -81.8583
1.6806 | -81.8558
1.6809 | -81.8526
1.6813 | -81.8497
1.6817 | -81.8453
1.682 | -81.8429
1.6824 | -81.8441
1.6827 | -81.8452
1.6831 | -81.8434
1.6835 | -81.8411
1.6838 | -81.8367
1.6842 | -81.8314
1.6845 | -81.8269
1.6849 | -81.821
1.6853 | -81.8165
1.6856 | -81.8098
1.686 | -81.8042
1.6864 | -81.7995
1.6867 | -81.7954
1.6871 | -81.7935
1.6874 | -81.7896
1.6878 | -81.7862
1.6882 | -81.7832
1.6885 | -81.7819
1.6889 | -81.7785
1.6892 | -81.7735
1.6896 | -81.7678
1.69 | -81.7626
1.6903 | -81.7582
1.6907 | -81.754
1.6911 | -81.751
1.6914 | -81.7459
1.6918 | -81.7418
1.6921 | -81.7374
1.6925 | -81.7317
1.6929 | -81.7264
1.6932 | -81.7209
1.6936 | -81.7163
1.694 | -81.7124
1.6943 | -81.7076
1.6947 | -81.7028
1.695 | -81.6992
1.6954 | -81.698
1.6958 | -81.6975
1.6961 | -81.6947
1.6965 | -81.69
1.6968 | -81.6835
1.6972 | -81.6787
1.6976 | -81.6741
1.6979 | -81.6686
1.6983 | -81.6606
1.6987 | -81.6543
1.699 | -81.6479
1.6994 | -81.6432
1.6997 | -81.6387
1.7001 | -81.6336
1.7005 | -81.6291
1.7008 | -81.6255
1.7012 | -81.6215
1.7015 | -81.6202
1.7019 | -81.6182
1.7023 | -81.6173
1.7026 | -81.6156
1.703 | -81.6133
1.7034 | -81.6085
1.7037 | -81.6036
1.7041 | -81.5981
1.7044 | -81.5951
1.7048 | -81.5928
1.7052 | -81.5879
1.7055 | -81.5831
1.7059 | -81.5801
1.7063 | -81.5745
1.7066 | -81.5694
1.707 | -81.5652
1.7073 | -81.5605
1.7077 | -81.5585
1.7081 | -81.5554
1.7084 | -81.5496
1.7088 | -81.5425
1.7091 | -81.5359
1.7095 | -81.5298
1.7099 | -81.5253
1.7102 | -81.5197
1.7106 | -81.516
1.711 | -81.5119
1.7113 | -81.5087
1.7117 | -81.5088
1.712 | -81.5043
1.7124 | -81.498
1.7128 | -81.4949
1.7131 | -81.4909
1.7135 | -81.4884
1.7139 | -81.4875
1.7142 | -81.4856
1.7146 | -81.4838
1.7149 | -81.4829
1.7153 | -81.4795
1.7157 | -81.4768
1.716 | -81.4748
1.7164 | -81.473
1.7167 | -81.4689
1.7171 | -81.4626
1.7175 | -81.457
1.7178 | -81.4523
1.7182 | -81.4504
1.7186 | -81.45
1.7189 | -81.4506
1.7193 | -81.451
1.7196 | -81.4524
1.72 | -81.4523
1.7204 | -81.4464
1.7207 | -81.4409
1.7211 | -81.4397
1.7214 | -81.4385
1.7218 | -81.4356
1.7222 | -81.4342
1.7225 | -81.4324
1.7229 | -81.4329
1.7233 | -81.4333
1.7236 | -81.4331
1.724 | -81.433
1.7243 | -81.4336
1.7247 | -81.4321
1.7251 | -81.4294
1.7254 | -81.4285
1.7258 | -81.4264
1.7262 | -81.4259
1.7265 | -81.4284
1.7269 | -81.4328
1.7272 | -81.4354
1.7276 | -81.4381
1.728 | -81.4417
1.7283 | -81.4403
1.7287 | -81.4372
1.729 | -81.433
1.7294 | -81.4314
1.7298 | -81.4293
1.7301 | -81.4269
1.7305 | -81.4294
1.7309 | -81.4304
1.7312 | -81.4299
1.7316 | -81.4272
1.7319 | -81.4251
1.7323 | -81.422
1.7327 | -81.4232
1.733 | -81.4213
1.7334 | -81.419
1.7338 | -81.4168
1.7341 | -81.4181
1.7345 | -81.4186
1.7348 | -81.4182
1.7352 | -81.4181
1.7356 | -81.4184
1.7359 | -81.4176
1.7363 | -81.4154
1.7366 | -81.4154
1.737 | -81.4166
1.7374 | -81.4191
1.7377 | -81.4225
1.7381 | -81.423
1.7385 | -81.424
1.7388 | -81.4249
1.7392 | -81.4247
1.7395 | -81.4226
1.7399 | -81.4203
1.7403 | -81.4196
1.7406 | -81.42
1.741 | -81.4202
1.7413 | -81.4211
1.7417 | -81.4219
1.7421 | -81.4232
1.7424 | -81.4256
1.7428 | -81.4271
1.7432 | -81.4289
1.7435 | -81.4289
1.7439 | -81.4279
1.7442 | -81.4269
1.7446 | -81.4267
1.745 | -81.4266
1.7453 | -81.429
1.7457 | -81.4317
1.7461 | -81.4338
1.7464 | -81.4343
1.7468 | -81.4344
1.7471 | -81.4337
1.7475 | -81.4343
1.7479 | -81.4356
1.7482 | -81.436
1.7486 | -81.4357
1.7489 | -81.4337
1.7493 | -81.4308
1.7497 | -81.4307
1.75 | -81.4331
1.7504 | -81.4354
1.7508 | -81.4393
1.7511 | -81.446
1.7515 | -81.4478
1.7518 | -81.4472
1.7522 | -81.4473
1.7526 | -81.4465
1.7529 | -81.4471
1.7533 | -81.448
1.7536 | -81.4499
1.754 | -81.452
1.7544 | -81.4534
1.7547 | -81.4514
1.7551 | -81.4512
1.7555 | -81.451
1.7558 | -81.4511
1.7562 | -81.4514
1.7565 | -81.4497
1.7569 | -81.448
1.7573 | -81.4484
1.7576 | -81.4465
1.758 | -81.4438
1.7584 | -81.4441
1.7587 | -81.445
1.7591 | -81.4458
1.7594 | -81.445
1.7598 | -81.4431
1.7602 | -81.4364
1.7605 | -81.4347
1.7609 | -81.4315
1.7612 | -81.4295
1.7616 | -81.4296
1.762 | -81.4296
1.7623 | -81.4298
1.7627 | -81.4287
1.7631 | -81.4286
1.7634 | -81.4285
1.7638 | -81.4265
1.7641 | -81.4215
1.7645 | -81.4218
1.7649 | -81.425
1.7652 | -81.4286
1.7656 | -81.428
1.766 | -81.4281
1.7663 | -81.433
1.7667 | -81.436
1.767 | -81.4348
1.7674 | -81.4304
1.7678 | -81.4268
1.7681 | -81.4259
1.7685 | -81.4262
1.7688 | -81.4264
1.7692 | -81.427
1.7696 | -81.427
1.7699 | -81.4261
1.7703 | -81.4264
1.7707 | -81.4281
1.771 | -81.4277
1.7714 | -81.4329
1.7717 | -81.4343
1.7721 | -81.4346
1.7725 | -81.4349
1.7728 | -81.4363
1.7732 | -81.437
1.7735 | -81.4378
1.7739 | -81.4386
1.7743 | -81.4407
1.7746 | -81.442
1.775 | -81.4417
1.7754 | -81.442
1.7757 | -81.4428
1.7761 | -81.4434
1.7764 | -81.4435
1.7768 | -81.4416
1.7772 | -81.4425
1.7775 | -81.4441
1.7779 | -81.4446
1.7783 | -81.4431
1.7786 | -81.4397
1.779 | -81.4354
1.7793 | -81.4412
1.7797 | -81.4463
1.7801 | -81.4463
1.7804 | -81.4479
1.7808 | -81.4544
1.7811 | -81.4572
1.7815 | -81.4573
1.7819 | -81.4563
1.7822 | -81.4579
1.7826 | -81.4608
1.783 | -81.4636
1.7833 | -81.4645
1.7837 | -81.4653
1.784 | -81.4655
1.7844 | -81.4666
1.7848 | -81.4686
1.7851 | -81.4697
1.7855 | -81.4714
1.7858 | -81.4737
1.7862 | -81.4756
1.7866 | -81.4803
1.7869 | -81.4813
1.7873 | -81.4844
1.7877 | -81.4889
1.788 | -81.4919
1.7884 | -81.4941
1.7887 | -81.4959
1.7891 | -81.4969
1.7895 | -81.4986
1.7898 | -81.4997
1.7902 | -81.5018
1.7906 | -81.5021
1.7909 | -81.5046
1.7913 | -81.5071
1.7916 | -81.5093
1.792 | -81.5125
1.7924 | -81.5148
1.7927 | -81.5171
1.7931 | -81.5168
1.7934 | -81.5176
1.7938 | -81.519
1.7942 | -81.5206
1.7945 | -81.5241
1.7949 | -81.5247
1.7953 | -81.5238
1.7956 | -81.5203
1.796 | -81.5197
1.7963 | -81.5204
1.7967 | -81.5213
1.7971 | -81.5228
1.7974 | -81.5233
1.7978 | -81.5241
1.7982 | -81.526
1.7985 | -81.5278
1.7989 | -81.528
1.7992 | -81.5284
1.7996 | -81.5307
1.8 | -81.5353
1.8003 | -81.5374
1.8007 | -81.5392
1.801 | -81.5417
1.8014 | -81.544
1.8018 | -81.5438
1.8021 | -81.5436
1.8025 | -81.5452
1.8029 | -81.545
1.8032 | -81.546
1.8036 | -81.5472
1.8039 | -81.5448
1.8043 | -81.5439
1.8047 | -81.5447
1.805 | -81.5464
1.8054 | -81.5488
1.8057 | -81.5496
1.8061 | -81.5495
1.8065 | -81.5492
1.8068 | -81.5515
1.8072 | -81.554
1.8076 | -81.5571
1.8079 | -81.5584
1.8083 | -81.5575
1.8086 | -81.555
1.809 | -81.5552
1.8094 | -81.5582
1.8097 | -81.5585
1.8101 | -81.5597
1.8105 | -81.5608
1.8108 | -81.5589
1.8112 | -81.5582
1.8115 | -81.5575
1.8119 | -81.5561
1.8123 | -81.5554
1.8126 | -81.5545
1.813 | -81.5545
1.8133 | -81.5547
1.8137 | -81.5544
1.8141 | -81.5554
1.8144 | -81.5552
1.8148 | -81.5557
1.8152 | -81.5564
1.8155 | -81.5549
1.8159 | -81.5512
1.8162 | -81.5487
1.8166 | -81.5492
1.817 | -81.549
1.8173 | -81.5489
1.8177 | -81.5482
1.818 | -81.5481
1.8184 | -81.5486
1.8188 | -81.5476
1.8191 | -81.548
1.8195 | -81.5495
1.8199 | -81.5488
1.8202 | -81.5483
1.8206 | -81.549
1.8209 | -81.5501
1.8213 | -81.5481
1.8217 | -81.5454
1.822 | -81.5416
1.8224 | -81.5416
1.8228 | -81.5417
1.8231 | -81.5368
1.8235 | -81.5324
1.8238 | -81.5328
1.8242 | -81.5317
1.8246 | -81.5313
1.8249 | -81.5305
1.8253 | -81.5299
1.8256 | -81.5305
1.826 | -81.5291
1.8264 | -81.5274
1.8267 | -81.5249
1.8271 | -81.5239
1.8275 | -81.524
1.8278 | -81.5235
1.8282 | -81.5218
1.8285 | -81.5219
1.8289 | -81.5228
1.8293 | -81.5219
1.8296 | -81.518
1.83 | -81.5145
1.8304 | -81.5126
1.8307 | -81.5087
1.8311 | -81.5053
1.8314 | -81.504
1.8318 | -81.5039
1.8322 | -81.5045
1.8325 | -81.5046
1.8329 | -81.5051
1.8332 | -81.505
1.8336 | -81.5038
1.834 | -81.5015
1.8343 | -81.5001
1.8347 | -81.4981
1.8351 | -81.4942
1.8354 | -81.4931
1.8358 | -81.4942
1.8361 | -81.4955
1.8365 | -81.4942
1.8369 | -81.4931
1.8372 | -81.4947
1.8376 | -81.4935
1.8379 | -81.4918
1.8383 | -81.4894
1.8387 | -81.4883
1.839 | -81.488
1.8394 | -81.4858
1.8398 | -81.4851
1.8401 | -81.4844
1.8405 | -81.4833
1.8408 | -81.4854
1.8412 | -81.4891
1.8416 | -81.4916
1.8419 | -81.4921
1.8423 | -81.4914
1.8427 | -81.4898
1.843 | -81.4888
1.8434 | -81.4892
1.8437 | -81.4886
1.8441 | -81.4877
1.8445 | -81.4912
1.8448 | -81.4961
1.8452 | -81.498
1.8455 | -81.4989
1.8459 | -81.5008
1.8463 | -81.503
1.8466 | -81.5044
1.847 | -81.5056
1.8474 | -81.5107
1.8477 | -81.5113
1.8481 | -81.5102
1.8484 | -81.5099
1.8488 | -81.5109
1.8492 | -81.5105
1.8495 | -81.512
1.8499 | -81.5143
1.8503 | -81.5161
1.8506 | -81.5177
1.851 | -81.5177
1.8513 | -81.5185
1.8517 | -81.5209
1.8521 | -81.5237
1.8524 | -81.5266
1.8528 | -81.5289
1.8531 | -81.5341
1.8535 | -81.5396
1.8539 | -81.5457
1.8542 | -81.5477
1.8546 | -81.5484
1.855 | -81.5518
1.8553 | -81.5529
1.8557 | -81.555
1.856 | -81.5584
1.8564 | -81.5642
1.8568 | -81.571
1.8571 | -81.5768
1.8575 | -81.5807
1.8578 | -81.5904
1.8582 | -81.5948
1.8586 | -81.5967
1.8589 | -81.5975
1.8593 | -81.6001
1.8597 | -81.6013
1.86 | -81.605
1.8604 | -81.6097
1.8607 | -81.614
1.8611 | -81.6182
1.8615 | -81.6235
1.8618 | -81.6279
1.8622 | -81.6312
1.8626 | -81.6352
1.8629 | -81.6393
1.8633 | -81.6416
1.8636 | -81.6475
1.864 | -81.6525
1.8644 | -81.6546
1.8647 | -81.6593
1.8651 | -81.6679
1.8654 | -81.6761
1.8658 | -81.6799
1.8662 | -81.6834
1.8665 | -81.6885
1.8669 | -81.6932
1.8673 | -81.6998
1.8676 | -81.706
1.868 | -81.7112
1.8683 | -81.7184
1.8687 | -81.7238
1.8691 | -81.73
1.8694 | -81.7362
1.8698 | -81.7429
1.8701 | -81.7499
1.8705 | -81.7563
1.8709 | -81.7621
1.8712 | -81.7684
1.8716 | -81.7753
1.872 | -81.7815
1.8723 | -81.7862
1.8727 | -81.7935
1.873 | -81.7998
1.8734 | -81.8075
1.8738 | -81.815
1.8741 | -81.8223
1.8745 | -81.8291
1.8749 | -81.8309
1.8752 | -81.834
1.8756 | -81.8421
1.8759 | -81.8496
1.8763 | -81.8561
1.8767 | -81.8628
1.877 | -81.8671
1.8774 | -81.8696
1.8777 | -81.8732
1.8781 | -81.8795
1.8785 | -81.8856
1.8788 | -81.8853
1.8792 | -81.8912
1.8796 | -81.8983
1.8799 | -81.9043
1.8803 | -81.9093
1.8806 | -81.9142
1.881 | -81.9197
1.8814 | -81.9272
1.8817 | -81.9378
1.8821 | -81.9449
1.8825 | -81.9506
1.8828 | -81.9568
1.8832 | -81.9626
1.8835 | -81.9704
1.8839 | -81.9779
1.8843 | -81.9837
1.8846 | -81.9915
1.885 | -82.0006
1.8853 | -82.0054
1.8857 | -82.0099
1.8861 | -82.0149
1.8864 | -82.0177
1.8868 | -82.0229
1.8872 | -82.0349
1.8875 | -82.0424
1.8879 | -82.0501
1.8882 | -82.0564
1.8886 | -82.0613
1.889 | -82.0644
1.8893 | -82.0613
1.8897 | -82.0612
1.89 | -82.0656
1.8904 | -82.0693
1.8908 | -82.0728
1.8911 | -82.0786
1.8915 | -82.0871
1.8919 | -82.0911
1.8922 | -82.0903
1.8926 | -82.0878
1.8929 | -82.0871
1.8933 | -82.0905
1.8937 | -82.0946
1.894 | -82.0982
1.8944 | -82.1011
1.8948 | -82.1048
1.8951 | -82.1127
1.8955 | -82.1197
1.8958 | -82.1232
1.8962 | -82.1194
1.8966 | -82.115
1.8969 | -82.1122
1.8973 | -82.1107
1.8976 | -82.1131
1.898 | -82.1141
1.8984 | -82.115
1.8987 | -82.1163
1.8991 | -82.1179
1.8995 | -82.1175
1.8998 | -82.1131
1.9002 | -82.1133
1.9005 | -82.1131
1.9009 | -82.113
1.9013 | -82.1093
1.9016 | -82.1069
1.902 | -82.107
1.9023 | -82.1068
1.9027 | -82.1054
1.9031 | -82.1035
1.9034 | -82.1043
1.9038 | -82.1043
1.9042 | -82.1028
1.9045 | -82.1017
1.9049 | -82.1011
1.9052 | -82.1006
1.9056 | -82.0997
1.906 | -82.0964
1.9063 | -82.0899
1.9067 | -82.0912
1.9071 | -82.0946
1.9074 | -82.0974
1.9078 | -82.0954
1.9081 | -82.0929
1.9085 | -82.0909
1.9089 | -82.0877
1.9092 | -82.0835
1.9096 | -82.081
1.9099 | -82.0705
1.9103 | -82.067
1.9107 | -82.0652
1.911 | -82.0596
1.9114 | -82.0557
1.9118 | -82.05
1.9121 | -82.0422
1.9125 | -82.0365
1.9128 | -82.0324
1.9132 | -82.0279
1.9136 | -82.022
1.9139 | -82.013
1.9143 | -82.0061
1.9147 | -82.001
1.915 | -81.9954
1.9154 | -81.9895
1.9157 | -81.9827
1.9161 | -81.9783
1.9165 | -81.9748
1.9168 | -81.9706
1.9172 | -81.9664
1.9175 | -81.9593
1.9179 | -81.9515
1.9183 | -81.9451
1.9186 | -81.9395
1.919 | -81.9358
1.9194 | -81.9326
1.9197 | -81.9305
1.9201 | -81.9274
1.9204 | -81.9185
1.9208 | -81.9124
1.9212 | -81.9037
1.9215 | -81.8971
1.9219 | -81.8928
1.9222 | -81.8878
1.9226 | -81.8892
1.923 | -81.8908
1.9233 | -81.8833
1.9237 | -81.8765
1.9241 | -81.868
1.9244 | -81.8594
1.9248 | -81.8497
1.9251 | -81.846
1.9255 | -81.8449
1.9259 | -81.8383
1.9262 | -81.8298
1.9266 | -81.8217
1.927 | -81.8161
1.9273 | -81.8102
1.9277 | -81.8016
1.928 | -81.7931
1.9284 | -81.7842
1.9288 | -81.7789
1.9291 | -81.7728
1.9295 | -81.7635
1.9298 | -81.7544
1.9302 | -81.7461
1.9306 | -81.7366
1.9309 | -81.7303
1.9313 | -81.7226
1.9317 | -81.7151
1.932 | -81.7081
1.9324 | -81.6977
1.9327 | -81.6854
1.9331 | -81.6784
1.9335 | -81.671
1.9338 | -81.6628
1.9342 | -81.6564
1.9345 | -81.6494
1.9349 | -81.6421
1.9353 | -81.6336
1.9356 | -81.6239
1.936 | -81.6149
1.9364 | -81.6075
1.9367 | -81.602
1.9371 | -81.5961
1.9374 | -81.5879
1.9378 | -81.5778
1.9382 | -81.5666
1.9385 | -81.5571
1.9389 | -81.5506
1.9393 | -81.5462
1.9396 | -81.5397
1.94 | -81.5325
1.9403 | -81.5263
1.9407 | -81.5194
1.9411 | -81.5131
1.9414 | -81.5065
1.9418 | -81.5005
1.9421 | -81.4945
1.9425 | -81.4895
1.9429 | -81.4842
1.9432 | -81.477
1.9436 | -81.4689
1.944 | -81.4614
1.9443 | -81.4542
1.9447 | -81.4473
1.945 | -81.4406
1.9454 | -81.4351
1.9458 | -81.429
1.9461 | -81.4225
1.9465 | -81.4185
1.9469 | -81.4163
1.9472 | -81.4117
1.9476 | -81.406
1.9479 | -81.3992
1.9483 | -81.394
1.9487 | -81.3887
1.949 | -81.3835
1.9494 | -81.3768
1.9497 | -81.3705
1.9501 | -81.3652
1.9505 | -81.3588
1.9508 | -81.3537
1.9512 | -81.348
1.9516 | -81.3425
1.9519 | -81.3411
1.9523 | -81.3376
1.9526 | -81.3352
1.953 | -81.3321
1.9534 | -81.3284
1.9537 | -81.3259
1.9541 | -81.3216
1.9544 | -81.3173
1.9548 | -81.3168
1.9552 | -81.3147
1.9555 | -81.3093
1.9559 | -81.3057
1.9563 | -81.3016
1.9566 | -81.2973
1.957 | -81.2919
1.9573 | -81.2864
1.9577 | -81.2818
1.9581 | -81.2775
1.9584 | -81.2721
1.9588 | -81.2695
1.9592 | -81.2672
1.9595 | -81.2641
1.9599 | -81.259
1.9602 | -81.2548
1.9606 | -81.2499
1.961 | -81.2466
1.9613 | -81.2459
1.9617 | -81.2478
1.962 | -81.2476
1.9624 | -81.2499
1.9628 | -81.2476
1.9631 | -81.2472
1.9635 | -81.2464
1.9639 | -81.2448
1.9642 | -81.2416
1.9646 | -81.2384
1.9649 | -81.2374
1.9653 | -81.2361
1.9657 | -81.2347
1.966 | -81.2341
1.9664 | -81.2309
1.9667 | -81.2268
1.9671 | -81.2244
1.9675 | -81.2236
1.9678 | -81.2224
1.9682 | -81.2215
1.9686 | -81.2203
1.9689 | -81.2165
1.9693 | -81.2144
1.9696 | -81.2138
1.97 | -81.2131
1.9704 | -81.2134
1.9707 | -81.2142
1.9711 | -81.2151
1.9715 | -81.2148
1.9718 | -81.2151
1.9722 | -81.2169
1.9725 | -81.2171
1.9729 | -81.2184
1.9733 | -81.219
1.9736 | -81.221
1.974 | -81.2202
1.9743 | -81.2201
1.9747 | -81.2186
1.9751 | -81.2191
1.9754 | -81.2201
1.9758 | -81.2222
1.9762 | -81.2238
1.9765 | -81.2223
1.9769 | -81.2246
1.9772 | -81.2245
1.9776 | -81.2264
1.978 | -81.2281
1.9783 | -81.2286
1.9787 | -81.2288
1.9791 | -81.2307
1.9794 | -81.2328
1.9798 | -81.2341
1.9801 | -81.2277
1.9805 | -81.2247
1.9809 | -81.2271
1.9812 | -81.2298
1.9816 | -81.2328
1.9819 | -81.2342
1.9823 | -81.2351
1.9827 | -81.2375
1.983 | -81.2423
1.9834 | -81.2453
1.9838 | -81.2477
1.9841 | -81.251
1.9845 | -81.254
1.9848 | -81.2575
1.9852 | -81.2606
1.9856 | -81.2627
1.9859 | -81.2655
1.9863 | -81.2684
1.9867 | -81.2704
1.987 | -81.2712
1.9874 | -81.2727
1.9877 | -81.2749
1.9881 | -81.2782
1.9885 | -81.2829
1.9888 | -81.2783
1.9892 | -81.2727
1.9895 | -81.2719
1.9899 | -81.273
1.9903 | -81.2758
1.9906 | -81.2784
1.991 | -81.2813
1.9914 | -81.2846
1.9917 | -81.2825
1.9921 | -81.2788
1.9924 | -81.2805
1.9928 | -81.2853
1.9932 | -81.2926
1.9935 | -81.2954
1.9939 | -81.2999
1.9942 | -81.3037
1.9946 | -81.3018
1.995 | -81.3037
1.9953 | -81.3057
1.9957 | -81.3077
1.9961 | -81.3108
1.9964 | -81.315
1.9968 | -81.319
1.9971 | -81.3221
1.9975 | -81.3251
1.9979 | -81.3295
1.9982 | -81.333
1.9986 | -81.3361
1.999 | -81.3387
1.9993 | -81.3413
1.9997 | -81.3421
2 | -81.3404
2.0004 | -81.3394
2.0008 | -81.3408
2.0011 | -81.3434
2.0015 | -81.3453
2.0018 | -81.3468
2.0022 | -81.348
2.0026 | -81.3503
2.0029 | -81.3512
2.0033 | -81.3535
2.0037 | -81.3547
2.004 | -81.3552
2.0044 | -81.3553
2.0047 | -81.3567
2.0051 | -81.3579
2.0055 | -81.3592
2.0058 | -81.36
2.0062 | -81.36
2.0065 | -81.3625
2.0069 | -81.3634
2.0073 | -81.3652
2.0076 | -81.3676
2.008 | -81.3686
2.0084 | -81.3708
2.0087 | -81.3725
2.0091 | -81.372
2.0094 | -81.3689
2.0098 | -81.37
2.0102 | -81.375
2.0105 | -81.3789
2.0109 | -81.3798
2.0113 | -81.3778
2.0116 | -81.3769
2.012 | -81.3763
2.0123 | -81.3786
2.0127 | -81.3817
2.0131 | -81.3817
2.0134 | -81.3824
2.0138 | -81.3831
2.0141 | -81.3859
2.0145 | -81.3856
2.0149 | -81.3888
2.0152 | -81.3893
2.0156 | -81.3888
2.016 | -81.3889
2.0163 | -81.389
2.0167 | -81.3895
2.017 | -81.3891
2.0174 | -81.388
2.0178 | -81.387
2.0181 | -81.3863
2.0185 | -81.3863
2.0189 | -81.3846
2.0192 | -81.385
2.0196 | -81.3837
2.0199 | -81.3833
2.0203 | -81.383
2.0207 | -81.3827
2.021 | -81.3809
2.0214 | -81.3812
2.0217 | -81.3806
2.0221 | -81.3806
2.0225 | -81.3821
2.0228 | -81.3821
2.0232 | -81.3778
2.0236 | -81.3762
2.0239 | -81.3757
2.0243 | -81.3751
2.0246 | -81.3741
2.025 | -81.3735
2.0254 | -81.3732
2.0257 | -81.3738
2.0261 | -81.3761
2.0264 | -81.3787
2.0268 | -81.378
2.0272 | -81.3781
2.0275 | -81.3762
2.0279 | -81.3747
2.0283 | -81.3737
2.0286 | -81.3727
2.029 | -81.3732
2.0293 | -81.3766
2.0297 | -81.3777
2.0301 | -81.3752
2.0304 | -81.3724
2.0308 | -81.3638
2.0312 | -81.352
2.0315 | -81.3458
2.0319 | -81.3408
2.0322 | -81.3381
2.0326 | -81.3378
2.033 | -81.3407
2.0333 | -81.3412
2.0337 | -81.3377
2.034 | -81.3336
2.0344 | -81.3285
2.0348 | -81.3263
2.0351 | -81.3236
2.0355 | -81.3202
2.0359 | -81.3177
2.0362 | -81.3136
2.0366 | -81.3115
2.0369 | -81.3105
2.0373 | -81.3089
2.0377 | -81.3065
2.038 | -81.3046
2.0384 | -81.3025
2.0387 | -81.3032
2.0391 | -81.3029
2.0395 | -81.3017
2.0398 | -81.2992
2.0402 | -81.2977
2.0406 | -81.296
2.0409 | -81.2927
2.0413 | -81.2905
2.0416 | -81.2892
2.042 | -81.2853
2.0424 | -81.2828
2.0427 | -81.2831
2.0431 | -81.2816
2.0435 | -81.279
2.0438 | -81.2803
2.0442 | -81.2779
2.0445 | -81.2748
2.0449 | -81.2726
2.0453 | -81.2717
2.0456 | -81.2701
2.046 | -81.2709
2.0463 | -81.2682
2.0467 | -81.2659
2.0471 | -81.2644
2.0474 | -81.2631
2.0478 | -81.263
2.0482 | -81.2652
2.0485 | -81.2663
2.0489 | -81.2672
2.0492 | -81.2677
2.0496 | -81.2656
2.05 | -81.264
2.0503 | -81.2642
2.0507 | -81.2636
2.0511 | -81.2609
2.0514 | -81.2548
2.0518 | -81.2542
2.0521 | -81.2572
2.0525 | -81.2568
2.0529 | -81.2556
2.0532 | -81.2554
2.0536 | -81.2563
2.0539 | -81.2541
2.0543 | -81.2499
2.0547 | -81.2478
2.055 | -81.2463
2.0554 | -81.2446
2.0558 | -81.244
2.0561 | -81.2455
2.0565 | -81.2459
2.0568 | -81.2454
2.0572 | -81.2443
2.0576 | -81.2439
2.0579 | -81.2427
2.0583 | -81.2405
2.0586 | -81.2391
2.059 | -81.2396
2.0594 | -81.2353
2.0597 | -81.2326
2.0601 | -81.2345
2.0605 | -81.2344
2.0608 | -81.2303
2.0612 | -81.2275
2.0615 | -81.2269
2.0619 | -81.2264
2.0623 | -81.2262
2.0626 | -81.2257
2.063 | -81.227
2.0634 | -81.2292
2.0637 | -81.231
2.0641 | -81.2295
2.0644 | -81.2301
2.0648 | -81.2303
2.0652 | -81.2318
2.0655 | -81.2338
2.0659 | -81.2341
2.0662 | -81.2328
2.0666 | -81.2348
2.067 | -81.2354
2.0673 | -81.2363
2.0677 | -81.2374
2.0681 | -81.2371
2.0684 | -81.2366
2.0688 | -81.2362
2.0691 | -81.2353
2.0695 | -81.2356
2.0699 | -81.238
2.0702 | -81.2388
2.0706 | -81.2388
2.0709 | -81.2396
2.0713 | -81.2407
2.0717 | -81.2418
2.072 | -81.2451
2.0724 | -81.2494
2.0728 | -81.25
2.0731 | -81.2514
2.0735 | -81.2537
2.0738 | -81.2562
2.0742 | -81.2576
2.0746 | -81.2583
2.0749 | -81.259
2.0753 | -81.2594
2.0757 | -81.2605
2.076 | -81.2593
2.0764 | -81.2586
2.0767 | -81.2593
2.0771 | -81.2612
2.0775 | -81.2603
2.0778 | -81.2585
2.0782 | -81.2604
2.0785 | -81.2622
2.0789 | -81.2615
2.0793 | -81.2613
2.0796 | -81.2609
2.08 | -81.2603
2.0804 | -81.2602
2.0807 | -81.26
2.0811 | -81.2599
2.0814 | -81.2609
2.0818 | -81.2631
2.0822 | -81.2641
2.0825 | -81.2632
2.0829 | -81.2624
2.0833 | -81.2626
2.0836 | -81.264
2.084 | -81.2664
2.0843 | -81.2665
2.0847 | -81.2668
2.0851 | -81.2648
2.0854 | -81.2656
2.0858 | -81.2654
2.0861 | -81.2664
2.0865 | -81.2673
2.0869 | -81.2673
2.0872 | -81.2668
2.0876 | -81.2683
2.088 | -81.2681
2.0883 | -81.2692
2.0887 | -81.2706
2.089 | -81.2705
2.0894 | -81.2707
2.0898 | -81.2697
2.0901 | -81.2686
2.0905 | -81.2695
2.0908 | -81.2698
2.0912 | -81.271
2.0916 | -81.271
2.0919 | -81.272
2.0923 | -81.2735
2.0927 | -81.275
2.093 | -81.2752
2.0934 | -81.2754
2.0937 | -81.2757
2.0941 | -81.2774
2.0945 | -81.2796
2.0948 | -81.2806
2.0952 | -81.2807
2.0956 | -81.2801
2.0959 | -81.2799
2.0963 | -81.2815
2.0966 | -81.2832
2.097 | -81.2835
2.0974 | -81.283
2.0977 | -81.2843
2.0981 | -81.2775
2.0984 | -81.2795
2.0988 | -81.2806
2.0992 | -81.2827
2.0995 | -81.2841
2.0999 | -81.2865
2.1003 | -81.287
2.1006 | -81.2858
2.101 | -81.2856
2.1013 | -81.2877
2.1017 | -81.2899
2.1021 | -81.2909
2.1024 | -81.2913
2.1028 | -81.2911
2.1031 | -81.2917
2.1035 | -81.2916
2.1039 | -81.2915
2.1042 | -81.2922
2.1046 | -81.2903
2.105 | -81.2904
2.1053 | -81.2906
2.1057 | -81.2914
2.106 | -81.2931
2.1064 | -81.293
2.1068 | -81.2909
2.1071 | -81.2916
2.1075 | -81.2911
2.1079 | -81.2918
2.1082 | -81.2928
2.1086 | -81.2869
2.1089 | -81.2864
2.1093 | -81.2881
2.1097 | -81.2904
2.11 | -81.2913
2.1104 | -81.2916
2.1107 | -81.2918
2.1111 | -81.2937
2.1115 | -81.2959
2.1118 | -81.2963
2.1122 | -81.2961
2.1126 | -81.2962
2.1129 | -81.2954
2.1133 | -81.296
2.1136 | -81.2974
2.114 | -81.2982
2.1144 | -81.2994
2.1147 | -81.3011
2.1151 | -81.3021
2.1155 | -81.3024
2.1158 | -81.302
2.1162 | -81.303
2.1165 | -81.305
2.1169 | -81.3055
2.1173 | -81.3064
2.1176 | -81.3076
2.118 | -81.3086
2.1183 | -81.31
2.1187 | -81.3116
2.1191 | -81.3154
2.1194 | -81.3194
2.1198 | -81.3207
2.1202 | -81.3248
2.1205 | -81.3258
2.1209 | -81.3262
2.1212 | -81.327
2.1216 | -81.3289
2.122 | -81.3305
2.1223 | -81.332
2.1227 | -81.3356
2.123 | -81.3448
2.1234 | -81.3514
2.1238 | -81.3521
2.1241 | -81.3534
2.1245 | -81.3552
2.1249 | -81.359
2.1252 | -81.3624
2.1256 | -81.3648
2.1259 | -81.367
2.1263 | -81.3684
2.1267 | -81.3693
2.127 | -81.3715
2.1274 | -81.3739
2.1278 | -81.3758
2.1281 | -81.3766
2.1285 | -81.3763
2.1288 | -81.3757
2.1292 | -81.3795
2.1296 | -81.3798
2.1299 | -81.3807
2.1303 | -81.3837
2.1306 | -81.3882
2.131 | -81.3918
2.1314 | -81.3945
2.1317 | -81.3956
2.1321 | -81.3983
2.1325 | -81.4084
2.1328 | -81.4186
2.1332 | -81.4244
2.1335 | -81.4274
2.1339 | -81.429
2.1343 | -81.4306
2.1346 | -81.4297
2.135 | -81.4294
2.1353 | -81.4362
2.1357 | -81.4433
2.1361 | -81.4465
2.1364 | -81.4498
2.1368 | -81.4509
2.1372 | -81.4536
2.1375 | -81.4556
2.1379 | -81.4589
2.1382 | -81.4609
2.1386 | -81.4666
2.139 | -81.4695
2.1393 | -81.4722
2.1397 | -81.4769
2.1401 | -81.4788
2.1404 | -81.48
2.1408 | -81.4829
2.1411 | -81.4861
2.1415 | -81.4889
2.1419 | -81.4907
2.1422 | -81.4939
2.1426 | -81.4973
2.1429 | -81.5009
2.1433 | -81.5057
2.1437 | -81.5093
2.144 | -81.5136
2.1444 | -81.5215
2.1448 | -81.5289
2.1451 | -81.5337
2.1455 | -81.5354
2.1458 | -81.5383
2.1462 | -81.5421
2.1466 | -81.5464
2.1469 | -81.5497
2.1473 | -81.5548
2.1477 | -81.5552
2.148 | -81.5565
2.1484 | -81.5604
2.1487 | -81.5648
2.1491 | -81.5694
2.1495 | -81.5726
2.1498 | -81.5758
2.1502 | -81.5797
2.1505 | -81.5804
2.1509 | -81.5841
2.1513 | -81.5884
2.1516 | -81.5878
2.152 | -81.5877
2.1524 | -81.5908
2.1527 | -81.5944
2.1531 | -81.5982
2.1534 | -81.5999
2.1538 | -81.6017
2.1542 | -81.6067
2.1545 | -81.613
2.1549 | -81.6149
2.1552 | -81.6179
2.1556 | -81.6224
2.156 | -81.6274
2.1563 | -81.6336
2.1567 | -81.6378
2.1571 | -81.6411
2.1574 | -81.6421
2.1578 | -81.6455
2.1581 | -81.6514
2.1585 | -81.653
2.1589 | -81.6556
2.1592 | -81.6591
2.1596 | -81.6628
2.16 | -81.6652
2.1603 | -81.6669
2.1607 | -81.6688
2.161 | -81.6708
2.1614 | -81.6735
2.1618 | -81.6726
2.1621 | -81.6725
2.1625 | -81.6745
2.1628 | -81.6774
2.1632 | -81.6814
2.1636 | -81.6851
2.1639 | -81.6872
2.1643 | -81.6893
2.1647 | -81.694
2.165 | -81.6958
2.1654 | -81.6966
2.1657 | -81.6963
2.1661 | -81.6959
2.1665 | -81.6976
2.1668 | -81.7014
2.1672 | -81.7025
2.1675 | -81.7048
2.1679 | -81.7038
2.1683 | -81.7011
2.1686 | -81.7016
2.169 | -81.7073
2.1694 | -81.7101
2.1697 | -81.7111
2.1701 | -81.7113
2.1704 | -81.7092
2.1708 | -81.7054
2.1712 | -81.7016
2.1715 | -81.7021
2.1719 | -81.7041
2.1723 | -81.7053
2.1726 | -81.7039
2.173 | -81.7035
2.1733 | -81.7003
2.1737 | -81.6967
2.1741 | -81.6982
2.1744 | -81.6991
2.1748 | -81.6989
2.1751 | -81.6997
2.1755 | -81.7012
2.1759 | -81.7018
2.1762 | -81.7018
2.1766 | -81.7014
2.177 | -81.7003
2.1773 | -81.7056
2.1777 | -81.7118
2.178 | -81.713
2.1784 | -81.7116
2.1788 | -81.7091
2.1791 | -81.7073
2.1795 | -81.7019
2.1799 | -81.7003
2.1802 | -81.6976
2.1806 | -81.6954
2.1809 | -81.6941
2.1813 | -81.6919
2.1817 | -81.6846
2.182 | -81.6805
2.1824 | -81.6796
2.1827 | -81.6771
2.1831 | -81.6716
2.1835 | -81.6682
2.1838 | -81.6636
2.1842 | -81.6599
2.1846 | -81.6563
2.1849 | -81.6527
2.1853 | -81.646
2.1856 | -81.6392
2.186 | -81.6353
2.1864 | -81.6313
2.1867 | -81.6274
2.1871 | -81.6251
2.1874 | -81.6208
2.1878 | -81.6167
2.1882 | -81.6137
2.1885 | -81.6093
2.1889 | -81.6018
2.1893 | -81.599
2.1896 | -81.5954
2.19 | -81.5882
2.1903 | -81.5844
2.1907 | -81.5806
2.1911 | -81.575
2.1914 | -81.569
2.1918 | -81.565
2.1922 | -81.5555
2.1925 | -81.5512
2.1929 | -81.5458
2.1932 | -81.5402
2.1936 | -81.531
2.194 | -81.5248
2.1943 | -81.5201
2.1947 | -81.5153
2.195 | -81.5108
2.1954 | -81.5038
2.1958 | -81.4978
2.1961 | -81.4915
2.1965 | -81.4848
2.1969 | -81.4785
2.1972 | -81.4714
2.1976 | -81.4644
2.1979 | -81.4598
2.1983 | -81.4558
2.1987 | -81.4511
2.199 | -81.4463
2.1994 | -81.4412
2.1997 | -81.4353
2.2001 | -81.4286
2.2005 | -81.422
2.2008 | -81.414
2.2012 | -81.4051
2.2016 | -81.3985
2.2019 | -81.393
2.2023 | -81.383
2.2026 | -81.3787
2.203 | -81.3726
2.2034 | -81.3656
2.2037 | -81.3594
2.2041 | -81.3526
2.2045 | -81.3454
2.2048 | -81.3383
2.2052 | -81.331
2.2055 | -81.3248
2.2059 | -81.3173
2.2063 | -81.3081
2.2066 | -81.3013
2.207 | -81.2957
2.2073 | -81.2898
2.2077 | -81.2821
2.2081 | -81.2758
2.2084 | -81.2694
2.2088 | -81.2605
2.2092 | -81.2512
2.2095 | -81.2452
2.2099 | -81.2394
2.2102 | -81.231
2.2106 | -81.223
2.211 | -81.2152
2.2113 | -81.2067
2.2117 | -81.1989
2.2121 | -81.1909
2.2124 | -81.1846
2.2128 | -81.1762
2.2131 | -81.1703
2.2135 | -81.1622
2.2139 | -81.1539
2.2142 | -81.1459
2.2146 | -81.1378
2.2149 | -81.1294
2.2153 | -81.1195
2.2157 | -81.1081
2.216 | -81.0991
2.2164 | -81.0917
2.2168 | -81.084
2.2171 | -81.0765
2.2175 | -81.0691
2.2178 | -81.062
2.2182 | -81.0542
2.2186 | -81.0465
2.2189 | -81.0381
2.2193 | -81.0308
2.2196 | -81.0229
2.22 | -81.0156
2.2204 | -81.0074
2.2207 | -80.9989
2.2211 | -80.991
2.2215 | -80.9848
2.2218 | -80.977
2.2222 | -80.9704
2.2225 | -80.9642
2.2229 | -80.9576
2.2233 | -80.9537
2.2236 | -80.9454
2.224 | -80.9393
2.2244 | -80.9312
2.2247 | -80.9216
2.2251 | -80.9153
2.2254 | -80.9093
2.2258 | -80.9033
2.2262 | -80.8981
2.2265 | -80.891
2.2269 | -80.8836
2.2272 | -80.8778
2.2276 | -80.873
2.228 | -80.8703
2.2283 | -80.8661
2.2287 | -80.8606
2.2291 | -80.8555
2.2294 | -80.8502
2.2298 | -80.8454
2.2301 | -80.8404
2.2305 | -80.8354
2.2309 | -80.8299
2.2312 | -80.8242
2.2316 | -80.8172
2.232 | -80.8121
2.2323 | -80.8075
2.2327 | -80.8019
2.233 | -80.7922
2.2334 | -80.782
2.2338 | -80.7804
2.2341 | -80.775
2.2345 | -80.7707
2.2348 | -80.7667
2.2352 | -80.7612
2.2356 | -80.7572
2.2359 | -80.7514
2.2363 | -80.7469
2.2367 | -80.7427
2.237 | -80.7388
2.2374 | -80.7354
2.2377 | -80.7306
2.2381 | -80.7257
2.2385 | -80.7225
2.2388 | -80.7189
2.2392 | -80.7157
2.2395 | -80.7127
2.2399 | -80.7104
2.2403 | -80.7077
2.2406 | -80.7052
2.241 | -80.7031
2.2414 | -80.701
2.2417 | -80.6992
2.2421 | -80.6977
2.2424 | -80.6962
2.2428 | -80.6947
2.2432 | -80.6929
2.2435 | -80.6907
2.2439 | -80.6866
2.2443 | -80.6888
2.2446 | -80.6874
2.245 | -80.6864
2.2453 | -80.6841
2.2457 | -80.6825
2.2461 | -80.682
2.2464 | -80.6824
2.2468 | -80.6829
2.2471 | -80.6824
2.2475 | -80.6806
2.2479 | -80.6796
2.2482 | -80.679
2.2486 | -80.678
2.249 | -80.6779
2.2493 | -80.681
2.2497 | -80.6801
2.25 | -80.677
2.2504 | -80.6781
2.2508 | -80.6785
2.2511 | -80.6792
2.2515 | -80.6801
2.2519 | -80.6813
2.2522 | -80.6831
2.2526 | -80.6842
2.2529 | -80.6846
2.2533 | -80.686
2.2537 | -80.6889
2.254 | -80.6912
2.2544 | -80.6927
2.2547 | -80.6979
2.2551 | -80.6964
2.2555 | -80.6962
2.2558 | -80.6984
2.2562 | -80.7
2.2566 | -80.7028
2.2569 | -80.7056
2.2573 | -80.7073
2.2576 | -80.7078
2.258 | -80.7102
2.2584 | -80.7145
2.2587 | -80.7185
2.2591 | -80.7206
2.2594 | -80.7236
2.2598 | -80.7262
2.2602 | -80.731
2.2605 | -80.7373
2.2609 | -80.7442
2.2613 | -80.75
2.2616 | -80.7535
2.262 | -80.7558
2.2623 | -80.759
2.2627 | -80.762
2.2631 | -80.7651
2.2634 | -80.7686
2.2638 | -80.772
2.2642 | -80.7739
2.2645 | -80.7763
2.2649 | -80.7832
2.2652 | -80.7876
2.2656 | -80.7926
2.266 | -80.7967
2.2663 | -80.7995
2.2667 | -80.8027
2.267 | -80.8075
2.2674 | -80.8115
2.2678 | -80.8168
2.2681 | -80.8233
2.2685 | -80.8281
2.2689 | -80.8327
2.2692 | -80.8379
2.2696 | -80.8437
2.2699 | -80.8492
2.2703 | -80.8548
2.2707 | -80.8604
2.271 | -80.8653
2.2714 | -80.8695
2.2717 | -80.8731
2.2721 | -80.8772
2.2725 | -80.8814
2.2728 | -80.8851
2.2732 | -80.8908
2.2736 | -80.8962
2.2739 | -80.9016
2.2743 | -80.9077
2.2746 | -80.9142
2.275 | -80.9198
2.2754 | -80.9279
2.2757 | -80.9336
2.2761 | -80.9392
2.2765 | -80.9453
2.2768 | -80.9511
2.2772 | -80.9577
2.2775 | -80.9635
2.2779 | -80.9697
2.2783 | -80.9747
2.2786 | -80.981
2.279 | -80.9841
2.2793 | -80.9903
2.2797 | -80.9964
2.2801 | -81.0028
2.2804 | -81.0098
2.2808 | -81.0162
2.2812 | -81.0218
2.2815 | -81.0249
2.2819 | -81.0309
2.2822 | -81.037
2.2826 | -81.0425
2.283 | -81.0491
2.2833 | -81.0552
2.2837 | -81.0614
2.2841 | -81.0681
2.2844 | -81.0741
2.2848 | -81.0764
2.2851 | -81.0823
2.2855 | -81.0901
2.2859 | -81.0952
2.2862 | -81.1021
2.2866 | -81.1087
2.2869 | -81.1155
2.2873 | -81.1208
2.2877 | -81.126
2.288 | -81.1307
2.2884 | -81.1347
2.2888 | -81.1428
2.2891 | -81.1488
2.2895 | -81.1527
2.2898 | -81.1591
2.2902 | -81.1643
2.2906 | -81.1696
2.2909 | -81.1745
2.2913 | -81.1789
2.2916 | -81.1843
2.292 | -81.1877
2.2924 | -81.1911
2.2927 | -81.1951
2.2931 | -81.1979
2.2935 | -81.2008
2.2938 | -81.206
2.2942 | -81.2116
2.2945 | -81.2178
2.2949 | -81.2244
2.2953 | -81.2301
2.2956 | -81.2349
2.296 | -81.238
2.2964 | -81.2384
2.2967 | -81.2445
2.2971 | -81.2476
2.2974 | -81.2498
2.2978 | -81.2551
2.2982 | -81.2602
2.2985 | -81.2651
2.2989 | -81.2703
2.2992 | -81.2743
2.2996 | -81.2787
2.3 | -81.2835
2.3003 | -81.2892
2.3007 | -81.2942
2.3011 | -81.2959
2.3014 | -81.2978
2.3018 | -81.3022
2.3021 | -81.3062
2.3025 | -81.313
2.3029 | -81.3199
2.3032 | -81.3242
2.3036 | -81.3281
2.3039 | -81.3311
2.3043 | -81.3345
2.3047 | -81.3377
2.305 | -81.3407
2.3054 | -81.3449
2.3058 | -81.3478
2.3061 | -81.3505
2.3065 | -81.3528
2.3068 | -81.3498
2.3072 | -81.3558
2.3076 | -81.3609
2.3079 | -81.3639
2.3083 | -81.3667
2.3087 | -81.3671
2.309 | -81.3657
2.3094 | -81.3662
2.3097 | -81.3684
2.3101 | -81.3716
2.3105 | -81.3752
2.3108 | -81.3781
2.3112 | -81.3792
2.3115 | -81.3802
2.3119 | -81.3807
2.3123 | -81.3812
2.3126 | -81.3797
2.313 | -81.3791
2.3134 | -81.3809
2.3137 | -81.3819
2.3141 | -81.3851
2.3144 | -81.3867
2.3148 | -81.388
2.3152 | -81.3889
2.3155 | -81.391
2.3159 | -81.392
2.3163 | -81.391
2.3166 | -81.3923
2.317 | -81.3944
2.3173 | -81.3926
2.3177 | -81.3928
2.3181 | -81.3965
2.3184 | -81.4002
2.3188 | -81.4018
2.3191 | -81.4023
2.3195 | -81.4001
2.3199 | -81.4001
2.3202 | -81.3996
2.3206 | -81.3976
2.321 | -81.3955
2.3213 | -81.3957
2.3217 | -81.3979
2.322 | -81.3999
2.3224 | -81.4022
2.3228 | -81.4018
2.3231 | -81.4013
2.3235 | -81.4026
2.3238 | -81.4032
2.3242 | -81.4044
2.3246 | -81.4081
2.3249 | -81.4114
2.3253 | -81.4105
2.3257 | -81.4087
2.326 | -81.408
2.3264 | -81.4085
2.3267 | -81.4076
2.3271 | -81.4089
2.3275 | -81.4097
2.3278 | -81.4077
2.3282 | -81.4077
2.3286 | -81.4065
2.3289 | -81.4035
2.3293 | -81.4017
2.3296 | -81.4013
2.33 | -81.3957
2.3304 | -81.3931
2.3307 | -81.3943
2.3311 | -81.3926
2.3314 | -81.3903
2.3318 | -81.3883
2.3322 | -81.3851
2.3325 | -81.3824
2.3329 | -81.3794
2.3333 | -81.3766
2.3336 | -81.375
2.334 | -81.3746
2.3343 | -81.3726
2.3347 | -81.3729
2.3351 | -81.3734
2.3354 | -81.3735
2.3358 | -81.3723
2.3361 | -81.37
2.3365 | -81.3697
2.3369 | -81.3666
2.3372 | -81.3672
2.3376 | -81.3673
2.338 | -81.3639
2.3383 | -81.3617
2.3387 | -81.3589
2.339 | -81.3561
2.3394 | -81.3549
2.3398 | -81.3531
2.3401 | -81.3494
2.3405 | -81.346
2.3409 | -81.3445
2.3412 | -81.3429
2.3416 | -81.3404
2.3419 | -81.3376
2.3423 | -81.3344
2.3427 | -81.3307
2.343 | -81.3286
2.3434 | -81.325
2.3437 | -81.3195
2.3441 | -81.3158
2.3445 | -81.3116
2.3448 | -81.3078
2.3452 | -81.3042
2.3456 | -81.3005
2.3459 | -81.2966
2.3463 | -81.2935
2.3466 | -81.2925
2.347 | -81.2897
2.3474 | -81.2863
2.3477 | -81.2822
2.3481 | -81.276
2.3485 | -81.2683
2.3488 | -81.2661
2.3492 | -81.2639
2.3495 | -81.2615
2.3499 | -81.2582
2.3503 | -81.2559
2.3506 | -81.2534
2.351 | -81.2506
2.3513 | -81.247
2.3517 | -81.2433
2.3521 | -81.2418
2.3524 | -81.2364
2.3528 | -81.2296
2.3532 | -81.2261
2.3535 | -81.223
2.3539 | -81.2203
2.3542 | -81.2171
2.3546 | -81.2135
2.355 | -81.2071
2.3553 | -81.2025
2.3557 | -81.2001
2.356 | -81.1961
2.3564 | -81.1913
2.3568 | -81.1871
2.3571 | -81.1823
2.3575 | -81.1784
2.3579 | -81.1772
2.3582 | -81.1737
2.3586 | -81.1704
2.3589 | -81.1652
2.3593 | -81.1609
2.3597 | -81.1573
2.36 | -81.1558
2.3604 | -81.1528
2.3608 | -81.1505
2.3611 | -81.1484
2.3615 | -81.1443
2.3618 | -81.1406
2.3622 | -81.137
2.3626 | -81.1334
2.3629 | -81.1307
2.3633 | -81.128
2.3636 | -81.1258
2.364 | -81.1235
2.3644 | -81.1212
2.3647 | -81.1177
2.3651 | -81.1115
2.3655 | -81.105
2.3658 | -81.1006
2.3662 | -81.0949
2.3665 | -81.0893
2.3669 | -81.0876
2.3673 | -81.087
2.3676 | -81.0853
2.368 | -81.0833
2.3683 | -81.0828
2.3687 | -81.0814
2.3691 | -81.079
2.3694 | -81.0797
2.3698 | -81.0778
2.3702 | -81.075
2.3705 | -81.0735
2.3709 | -81.0719
2.3712 | -81.0694
2.3716 | -81.0677
2.372 | -81.0639
2.3723 | -81.0606
2.3727 | -81.0579
2.3731 | -81.0554
2.3734 | -81.0538
2.3738 | -81.0533
2.3741 | -81.0524
2.3745 | -81.0502
2.3749 | -81.0465
2.3752 | -81.0471
2.3756 | -81.0475
2.3759 | -81.044
2.3763 | -81.0414
2.3767 | -81.0397
2.377 | -81.0378
2.3774 | -81.0366
2.3778 | -81.0351
2.3781 | -81.0354
2.3785 | -81.0361
2.3788 | -81.0356
2.3792 | -81.0341
2.3796 | -81.0339
2.3799 | -81.0366
2.3803 | -81.0346
2.3807 | -81.031
2.381 | -81.0275
2.3814 | -81.0233
2.3817 | -81.0201
2.3821 | -81.0199
2.3825 | -81.019
2.3828 | -81.0166
2.3832 | -81.0149
2.3835 | -81.0086
2.3839 | -81.0044
2.3843 | -81.0041
2.3846 | -81.0029
2.385 | -81.0042
2.3854 | -81.0072
2.3857 | -81.0061
2.3861 | -81.0048
2.3864 | -81.0048
2.3868 | -81.004
2.3872 | -81.0009
2.3875 | -80.9993
2.3879 | -80.999
2.3882 | -80.9988
2.3886 | -80.9974
2.389 | -80.9939
2.3893 | -80.9889
2.3897 | -80.9847
2.3901 | -80.984
2.3904 | -80.9852
2.3908 | -80.9836
2.3911 | -80.9817
2.3915 | -80.9782
2.3919 | -80.9774
2.3922 | -80.9779
2.3926 | -80.9789
2.393 | -80.9785
2.3933 | -80.9766
2.3937 | -80.9782
2.394 | -80.9828
2.3944 | -80.9879
2.3948 | -80.9888
2.3951 | -80.9872
2.3955 | -80.9872
2.3958 | -80.987
2.3962 | -80.9873
2.3966 | -80.9885
2.3969 | -80.9906
2.3973 | -80.9919
2.3977 | -80.9926
2.398 | -80.9925
2.3984 | -80.9917
2.3987 | -80.9913
2.3991 | -80.993
2.3995 | -80.9934
2.3998 | -80.9924
2.4002 | -80.9934
2.4005 | -80.9935
2.4009 | -80.995
2.4013 | -80.9951
2.4016 | -80.9924
2.402 | -80.9915
2.4024 | -80.991
2.4027 | -80.9909
2.4031 | -80.9911
2.4034 | -80.9891
2.4038 | -80.9876
2.4042 | -80.9865
2.4045 | -80.9854
2.4049 | -80.9848
2.4053 | -80.9848
2.4056 | -80.9875
2.406 | -80.989
2.4063 | -80.9889
2.4067 | -80.9892
2.4071 | -80.9894
2.4074 | -80.9907
2.4078 | -80.9924
2.4081 | -80.9927
2.4085 | -80.9924
2.4089 | -80.9914
2.4092 | -80.9898
2.4096 | -80.9907
2.41 | -80.9908
2.4103 | -80.9906
2.4107 | -80.9927
2.411 | -80.9932
2.4114 | -80.9913
2.4118 | -80.9909
2.4121 | -80.9937
2.4125 | -80.9968
2.4129 | -80.9964
2.4132 | -80.9957
2.4136 | -80.9959
2.4139 | -80.9959
2.4143 | -80.9964
2.4147 | -80.9969
2.415 | -80.997
2.4154 | -80.997
2.4157 | -80.9934
2.4161 | -80.9913
2.4165 | -80.9907
2.4168 | -80.9906
2.4172 | -80.9924
2.4176 | -80.9931
2.4179 | -80.9963
2.4183 | -80.9986
2.4186 | -80.9986
2.419 | -80.9968
2.4194 | -80.9978
2.4197 | -80.9963
2.4201 | -80.9976
2.4204 | -80.9996
2.4208 | -80.999
2.4212 | -80.9993
2.4215 | -80.9948
2.4219 | -80.9956
2.4223 | -80.9941
2.4226 | -80.9938
2.423 | -80.9945
2.4233 | -80.9944
2.4237 | -80.9946
2.4241 | -80.994
2.4244 | -80.9933
2.4248 | -80.9927
2.4252 | -80.9936
2.4255 | -80.9942
2.4259 | -80.994
2.4262 | -80.9937
2.4266 | -80.9954
2.427 | -80.9962
2.4273 | -80.999
2.4277 | -81.0011
2.428 | -81.0016
2.4284 | -81.0005
2.4288 | -80.9971
2.4291 | -80.9942
2.4295 | -80.9939
2.4299 | -80.9937
2.4302 | -80.9929
2.4306 | -80.9913
2.4309 | -80.9905
2.4313 | -80.9887
2.4317 | -80.9879
2.432 | -80.9858
2.4324 | -80.9861
2.4327 | -80.9836
2.4331 | -80.9833
2.4335 | -80.9823
2.4338 | -80.9838
2.4342 | -80.9839
2.4346 | -80.9828
2.4349 | -80.9822
2.4353 | -80.9828
2.4356 | -80.9845
2.436 | -80.9844
2.4364 | -80.985
2.4367 | -80.9858
2.4371 | -80.9853
2.4375 | -80.9858
2.4378 | -80.987
2.4382 | -80.9878
2.4385 | -80.9874
2.4389 | -80.985
2.4393 | -80.9837
2.4396 | -80.9824
2.44 | -80.9815
2.4403 | -80.9796
2.4407 | -80.9751
2.4411 | -80.9697
2.4414 | -80.9679
2.4418 | -80.9679
2.4422 | -80.9684
2.4425 | -80.9652
2.4429 | -80.9638
2.4432 | -80.9627
2.4436 | -80.965
2.444 | -80.965
2.4443 | -80.9642
2.4447 | -80.965
2.4451 | -80.9648
2.4454 | -80.9636
2.4458 | -80.9626
2.4461 | -80.9653
2.4465 | -80.9657
2.4469 | -80.9665
2.4472 | -80.9673
2.4476 | -80.962
2.4479 | -80.9595
2.4483 | -80.9594
2.4487 | -80.959
2.449 | -80.9588
2.4494 | -80.9613
2.4498 | -80.9609
2.4501 | -80.9607
2.4505 | -80.9598
2.4508 | -80.96
2.4512 | -80.9594
2.4516 | -80.9573
2.4519 | -80.9545
2.4523 | -80.9503
2.4526 | -80.9482
2.453 | -80.9483
2.4534 | -80.9488
2.4537 | -80.9483
2.4541 | -80.9485
2.4545 | -80.9473
2.4548 | -80.946
2.4552 | -80.9429
2.4555 | -80.9388
2.4559 | -80.9384
2.4563 | -80.9393
2.4566 | -80.9394
2.457 | -80.9389
2.4574 | -80.9399
2.4577 | -80.9417
2.4581 | -80.9417
2.4584 | -80.9413
2.4588 | -80.9424
2.4592 | -80.9413
2.4595 | -80.9398
2.4599 | -80.9347
2.4602 | -80.9322
2.4606 | -80.9307
2.461 | -80.9296
2.4613 | -80.9287
2.4617 | -80.9288
2.4621 | -80.9298
2.4624 | -80.929
2.4628 | -80.9297
2.4631 | -80.9307
2.4635 | -80.9278
2.4639 | -80.9264
2.4642 | -80.9253
2.4646 | -80.9228
2.465 | -80.9189
2.4653 | -80.9143
2.4657 | -80.9131
2.466 | -80.9115
2.4664 | -80.9098
2.4668 | -80.9093
2.4671 | -80.9111
2.4675 | -80.9122
2.4678 | -80.9114
2.4682 | -80.9105
2.4686 | -80.9073
2.4689 | -80.9052
2.4693 | -80.9047
2.4697 | -80.904
2.47 | -80.9036
2.4704 | -80.9034
2.4707 | -80.903
2.4711 | -80.9022
2.4715 | -80.9022
2.4718 | -80.904
2.4722 | -80.9022
2.4725 | -80.9008
2.4729 | -80.9019
2.4733 | -80.9015
2.4736 | -80.8977
2.474 | -80.8972
2.4744 | -80.8967
2.4747 | -80.8956
2.4751 | -80.8946
2.4754 | -80.8949
2.4758 | -80.898
2.4762 | -80.8996
2.4765 | -80.8986
2.4769 | -80.8976
2.4773 | -80.8961
2.4776 | -80.8949
2.478 | -80.8948
2.4783 | -80.8938
2.4787 | -80.8939
2.4791 | -80.8944
2.4794 | -80.8966
2.4798 | -80.8983
2.4801 | -80.8963
2.4805 | -80.8874
2.4809 | -80.882
2.4812 | -80.8808
2.4816 | -80.8793
2.482 | -80.8781
2.4823 | -80.8757
2.4827 | -80.874
2.483 | -80.8733
2.4834 | -80.8728
2.4838 | -80.871
2.4841 | -80.8703
2.4845 | -80.8693
2.4848 | -80.8678
2.4852 | -80.8673
2.4856 | -80.8655
2.4859 | -80.8635
2.4863 | -80.8614
2.4867 | -80.8592
2.487 | -80.8578
2.4874 | -80.8551
2.4877 | -80.8544
2.4881 | -80.8553
2.4885 | -80.8553
2.4888 | -80.8531
2.4892 | -80.8513
2.4896 | -80.8499
2.4899 | -80.8481
2.4903 | -80.8466
2.4906 | -80.8439
2.491 | -80.8407
2.4914 | -80.8361
2.4917 | -80.8327
2.4921 | -80.8284
2.4924 | -80.8234
2.4928 | -80.8204
2.4932 | -80.8175
2.4935 | -80.8115
2.4939 | -80.8049
2.4943 | -80.8044
2.4946 | -80.803
2.495 | -80.8022
2.4953 | -80.7961
2.4957 | -80.7927
2.4961 | -80.7907
2.4964 | -80.788
2.4968 | -80.7842
2.4972 | -80.7814
2.4975 | -80.7798
2.4979 | -80.7784
2.4982 | -80.7798
2.4986 | -80.781
2.499 | -80.7771
2.4993 | -80.7734
2.4997 | -80.7702
2.5 | -80.7667
2.5004 | -80.7657
2.5008 | -80.7641
2.5011 | -80.7616
2.5015 | -80.7604
2.5019 | -80.7588
2.5022 | -80.7564
2.5026 | -80.7542
2.5029 | -80.7499
2.5033 | -80.7464
2.5037 | -80.741
2.504 | -80.7348
2.5044 | -80.7301
2.5047 | -80.7271
2.5051 | -80.7237
2.5055 | -80.7219
2.5058 | -80.7192
2.5062 | -80.7167
2.5066 | -80.7146
2.5069 | -80.7105
2.5073 | -80.7081
2.5076 | -80.7058
2.508 | -80.7033
2.5084 | -80.7015
2.5087 | -80.6976
2.5091 | -80.6924
2.5095 | -80.6891
2.5098 | -80.687
2.5102 | -80.684
2.5105 | -80.6793
2.5109 | -80.6761
2.5113 | -80.6736
2.5116 | -80.6689
2.512 | -80.6594
2.5123 | -80.6553
2.5127 | -80.6539
2.5131 | -80.6515
2.5134 | -80.6495
2.5138 | -80.6464
2.5142 | -80.6401
2.5145 | -80.6334
2.5149 | -80.6284
2.5152 | -80.6259
2.5156 | -80.6233
2.516 | -80.6201
2.5163 | -80.6165
2.5167 | -80.6125
2.517 | -80.6111
2.5174 | -80.6099
2.5178 | -80.6088
2.5181 | -80.6072
2.5185 | -80.6025
2.5189 | -80.5973
2.5192 | -80.5939
2.5196 | -80.5895
2.5199 | -80.5864
2.5203 | -80.5828
2.5207 | -80.5793
2.521 | -80.5764
2.5214 | -80.5732
2.5218 | -80.5704
2.5221 | -80.5703
2.5225 | -80.5693
2.5228 | -80.5658
2.5232 | -80.5625
2.5236 | -80.5584
2.5239 | -80.5521
2.5243 | -80.5481
2.5246 | -80.5454
2.525 | -80.5397
2.5254 | -80.5349
2.5257 | -80.533
2.5261 | -80.5307
2.5265 | -80.5287
2.5268 | -80.5258
2.5272 | -80.5231
2.5275 | -80.521
2.5279 | -80.5175
2.5283 | -80.514
2.5286 | -80.5096
2.529 | -80.5052
2.5294 | -80.5016
2.5297 | -80.4993
2.5301 | -80.4961
2.5304 | -80.4919
2.5308 | -80.4874
2.5312 | -80.4837
2.5315 | -80.4802
2.5319 | -80.4778
2.5322 | -80.475
2.5326 | -80.4722
2.533 | -80.469
2.5333 | -80.4656
2.5337 | -80.4618
2.5341 | -80.4583
2.5344 | -80.4546
2.5348 | -80.4514
2.5351 | -80.4488
2.5355 | -80.4486
2.5359 | -80.4499
2.5362 | -80.4482
2.5366 | -80.447
2.5369 | -80.4463
2.5373 | -80.4457
2.5377 | -80.4428
2.538 | -80.4397
2.5384 | -80.4365
2.5388 | -80.4338
2.5391 | -80.4299
2.5395 | -80.4268
2.5398 | -80.4215
2.5402 | -80.4161
2.5406 | -80.4129
2.5409 | -80.4089
2.5413 | -80.4033
2.5417 | -80.3975
2.542 | -80.3933
2.5424 | -80.3889
2.5427 | -80.3857
2.5431 | -80.3844
2.5435 | -80.3833
2.5438 | -80.3812
2.5442 | -80.3782
2.5445 | -80.3748
2.5449 | -80.3722
2.5453 | -80.3692
2.5456 | -80.3668
2.546 | -80.3652
2.5464 | -80.3603
2.5467 | -80.3535
2.5471 | -80.3517
2.5474 | -80.3508
2.5478 | -80.3482
2.5482 | -80.3446
2.5485 | -80.3423
2.5489 | -80.3389
2.5492 | -80.3358
2.5496 | -80.333
2.55 | -80.33
2.5503 | -80.3276
2.5507 | -80.3242
2.5511 | -80.3213
2.5514 | -80.3187
2.5518 | -80.3156
2.5521 | -80.3139
2.5525 | -80.3129
2.5529 | -80.3057
2.5532 | -80.3015
2.5536 | -80.3004
2.554 | -80.2977
2.5543 | -80.2947
2.5547 | -80.29
2.555 | -80.2873
2.5554 | -80.288
2.5558 | -80.288
2.5561 | -80.2857
2.5565 | -80.2833
2.5568 | -80.2807
2.5572 | -80.2775
2.5576 | -80.2752
2.5579 | -80.2734
2.5583 | -80.2716
2.5587 | -80.269
2.559 | -80.266
2.5594 | -80.2651
2.5597 | -80.2633
2.5601 | -80.2602
2.5605 | -80.258
2.5608 | -80.2565
2.5612 | -80.2559
2.5616 | -80.2558
2.5619 | -80.256
2.5623 | -80.2534
2.5626 | -80.2507
2.563 | -80.247
2.5634 | -80.2433
2.5637 | -80.242
2.5641 | -80.2376
2.5644 | -80.2322
2.5648 | -80.2279
2.5652 | -80.2234
2.5655 | -80.2185
2.5659 | -80.2139
2.5663 | -80.2088
2.5666 | -80.2051
2.567 | -80.2015
2.5673 | -80.1977
2.5677 | -80.1915
2.5681 | -80.1834
2.5684 | -80.1777
2.5688 | -80.1736
2.5691 | -80.1688
2.5695 | -80.1632
2.5699 | -80.1571
2.5702 | -80.1538
2.5706 | -80.1513
2.571 | -80.1488
2.5713 | -80.1453
2.5717 | -80.1408
2.572 | -80.1369
2.5724 | -80.1339
2.5728 | -80.1313
2.5731 | -80.1277
2.5735 | -80.1229
2.5739 | -80.1192
2.5742 | -80.1151
2.5746 | -80.1119
2.5749 | -80.1103
2.5753 | -80.1082
2.5757 | -80.1033
2.576 | -80.0999
2.5764 | -80.0971
2.5767 | -80.0947
2.5771 | -80.0884
2.5775 | -80.0841
2.5778 | -80.082
2.5782 | -80.0805
2.5786 | -80.0782
2.5789 | -80.0753
2.5793 | -80.0729
2.5796 | -80.0699
2.58 | -80.066
2.5804 | -80.063
2.5807 | -80.0598
2.5811 | -80.057
2.5814 | -80.0525
2.5818 | -80.0486
2.5822 | -80.0462
2.5825 | -80.0448
2.5829 | -80.0442
2.5833 | -80.0424
2.5836 | -80.0406
2.584 | -80.0387
2.5843 | -80.0359
2.5847 | -80.0341
2.5851 | -80.0305
2.5854 | -80.0288
2.5858 | -80.0278
2.5862 | -80.0233
2.5865 | -80.0169
2.5869 | -80.0099
2.5872 | -80.0034
2.5876 | -79.9973
2.588 | -79.9898
2.5883 | -79.9798
2.5887 | -79.9694
2.589 | -79.9655
2.5894 | -79.9622
2.5898 | -79.9571
2.5901 | -79.9528
2.5905 | -79.9511
2.5909 | -79.9465
2.5912 | -79.9386
2.5916 | -79.9323
2.5919 | -79.9278
2.5923 | -79.924
2.5927 | -79.9195
2.593 | -79.9159
2.5934 | -79.9118
2.5938 | -79.9056
2.5941 | -79.9026
2.5945 | -79.8989
2.5948 | -79.8952
2.5952 | -79.8921
2.5956 | -79.8884
2.5959 | -79.8865
2.5963 | -79.8844
2.5966 | -79.8798
2.597 | -79.8763
2.5974 | -79.8722
2.5977 | -79.8682
2.5981 | -79.8638
2.5985 | -79.858
2.5988 | -79.8522
2.5992 | -79.8468
2.5995 | -79.8437
2.5999 | -79.8403
2.6003 | -79.8363
2.6006 | -79.832
2.601 | -79.8284
2.6014 | -79.8256
2.6017 | -79.82
2.6021 | -79.8145
2.6024 | -79.8094
2.6028 | -79.8032
2.6032 | -79.7988
2.6035 | -79.7955
2.6039 | -79.7923
2.6042 | -79.7889
2.6046 | -79.7839
2.605 | -79.7773
2.6053 | -79.7692
2.6057 | -79.7642
2.6061 | -79.7601
2.6064 | -79.7568
2.6068 | -79.753
2.6071 | -79.7484
2.6075 | -79.7421
2.6079 | -79.7366
2.6082 | -79.7315
2.6086 | -79.7248
2.6089 | -79.72
2.6093 | -79.7142
2.6097 | -79.7068
2.61 | -79.7011
2.6104 | -79.6959
2.6108 | -79.6911
2.6111 | -79.6855
2.6115 | -79.6799
2.6118 | -79.6751
2.6122 | -79.6694
2.6126 | -79.6636
2.6129 | -79.6571
2.6133 | -79.6501
2.6137 | -79.6449
2.614 | -79.6392
2.6144 | -79.6342
2.6147 | -79.6296
2.6151 | -79.6242
2.6155 | -79.6183
2.6158 | -79.6136
2.6162 | -79.6096
2.6165 | -79.6075
2.6169 | -79.6058
2.6173 | -79.6009
2.6176 | -79.594
2.618 | -79.587
2.6184 | -79.5804
2.6187 | -79.5735
2.6191 | -79.5641
2.6194 | -79.5576
2.6198 | -79.5512
2.6202 | -79.5459
2.6205 | -79.539
2.6209 | -79.5309
2.6212 | -79.5223
2.6216 | -79.5139
2.622 | -79.5066
2.6223 | -79.5014
2.6227 | -79.4965
2.6231 | -79.4915
2.6234 | -79.486
2.6238 | -79.4795
2.6241 | -79.4753
2.6245 | -79.4702
2.6249 | -79.4645
2.6252 | -79.4584
2.6256 | -79.4522
2.626 | -79.446
2.6263 | -79.4398
2.6267 | -79.4334
2.627 | -79.4269
2.6274 | -79.42
2.6278 | -79.4115
2.6281 | -79.4031
2.6285 | -79.3975
2.6288 | -79.394
2.6292 | -79.3894
2.6296 | -79.3836
2.6299 | -79.3776
2.6303 | -79.3675
2.6307 | -79.359
2.631 | -79.3524
2.6314 | -79.3429
2.6317 | -79.3342
2.6321 | -79.3271
2.6325 | -79.3206
2.6328 | -79.3158
2.6332 | -79.3094
2.6336 | -79.3028
2.6339 | -79.2967
2.6343 | -79.2906
2.6346 | -79.2841
2.635 | -79.2777
2.6354 | -79.2717
2.6357 | -79.2661
2.6361 | -79.2583
2.6364 | -79.2517
2.6368 | -79.246
2.6372 | -79.2418
2.6375 | -79.2352
2.6379 | -79.2259
2.6383 | -79.2163
2.6386 | -79.2109
2.639 | -79.205
2.6393 | -79.1989
2.6397 | -79.1926
2.6401 | -79.1864
2.6404 | -79.1804
2.6408 | -79.1766
2.6411 | -79.1714
2.6415 | -79.167
2.6419 | -79.161
2.6422 | -79.154
2.6426 | -79.1481
2.643 | -79.1441
2.6433 | -79.1407
2.6437 | -79.1353
2.644 | -79.1299
2.6444 | -79.1245
2.6448 | -79.1163
2.6451 | -79.1056
2.6455 | -79.0949
2.6459 | -79.0858
2.6462 | -79.0793
2.6466 | -79.0752
2.6469 | -79.0697
2.6473 | -79.0625
2.6477 | -79.0566
2.648 | -79.0506
2.6484 | -79.0436
2.6487 | -79.0358
2.6491 | -79.0277
2.6495 | -79.0211
2.6498 | -79.0156
2.6502 | -79.0122
2.6506 | -79.0085
2.6509 | -79.0027
2.6513 | -78.9991
2.6516 | -78.9946
2.652 | -78.9906
2.6524 | -78.9852
2.6527 | -78.98
2.6531 | -78.975
2.6534 | -78.97
2.6538 | -78.965
2.6542 | -78.96
2.6545 | -78.9541
2.6549 | -78.9482
2.6553 | -78.9432
2.6556 | -78.9385
2.656 | -78.9334
2.6563 | -78.9286
2.6567 | -78.9237
2.6571 | -78.9193
2.6574 | -78.9148
2.6578 | -78.9087
2.6582 | -78.9031
2.6585 | -78.8986
2.6589 | -78.8943
2.6592 | -78.8949
2.6596 | -78.893
2.66 | -78.8884
2.6603 | -78.8841
2.6607 | -78.88
2.661 | -78.8762
2.6614 | -78.8727
2.6618 | -78.8724
2.6621 | -78.8695
2.6625 | -78.8653
2.6629 | -78.8599
2.6632 | -78.8542
2.6636 | -78.8505
2.6639 | -78.8473
2.6643 | -78.8428
2.6647 | -78.8367
2.665 | -78.8337
2.6654 | -78.8299
2.6658 | -78.8267
2.6661 | -78.8244
2.6665 | -78.822
2.6668 | -78.8196
2.6672 | -78.8172
2.6676 | -78.8145
2.6679 | -78.8121
2.6683 | -78.8101
2.6686 | -78.8079
2.669 | -78.8035
2.6694 | -78.7996
2.6697 | -78.7977
2.6701 | -78.7961
2.6705 | -78.7937
2.6708 | -78.7923
2.6712 | -78.7911
2.6715 | -78.7895
2.6719 | -78.7885
2.6723 | -78.7903
2.6726 | -78.789
2.673 | -78.7879
2.6733 | -78.7879
2.6737 | -78.7881
2.6741 | -78.7864
2.6744 | -78.7847
2.6748 | -78.7825
2.6752 | -78.7802
2.6755 | -78.7811
2.6759 | -78.7796
2.6762 | -78.7776
2.6766 | -78.7752
2.677 | -78.7737
2.6773 | -78.7732
2.6777 | -78.7721
2.6781 | -78.7709
2.6784 | -78.77
2.6788 | -78.7691
2.6791 | -78.7691
2.6795 | -78.7666
2.6799 | -78.7637
2.6802 | -78.763
2.6806 | -78.7618
2.6809 | -78.7598
2.6813 | -78.7581
2.6817 | -78.7578
2.682 | -78.7555
2.6824 | -78.7524
2.6828 | -78.7525
2.6831 | -78.7522
2.6835 | -78.7527
2.6838 | -78.7516
2.6842 | -78.7522
2.6846 | -78.7528
2.6849 | -78.753
2.6853 | -78.7537
2.6856 | -78.7542
2.686 | -78.7548
2.6864 | -78.7548
2.6867 | -78.7552
2.6871 | -78.7545
2.6875 | -78.7534
2.6878 | -78.7507
2.6882 | -78.7489
2.6885 | -78.7475
2.6889 | -78.7474
2.6893 | -78.7471
2.6896 | -78.7465
2.69 | -78.7481
2.6904 | -78.7494
2.6907 | -78.7501
2.6911 | -78.7509
2.6914 | -78.7524
2.6918 | -78.7536
2.6922 | -78.754
2.6925 | -78.7562
2.6929 | -78.7572
2.6932 | -78.7613
2.6936 | -78.7636
2.694 | -78.7633
2.6943 | -78.7641
2.6947 | -78.7639
2.6951 | -78.7636
2.6954 | -78.7669
2.6958 | -78.7706
2.6961 | -78.7729
2.6965 | -78.7739
2.6969 | -78.7741
2.6972 | -78.7745
2.6976 | -78.7757
2.698 | -78.777
2.6983 | -78.7792
2.6987 | -78.7804
2.699 | -78.7819
2.6994 | -78.7839
2.6998 | -78.7854
2.7001 | -78.7875
2.7005 | -78.7905
2.7008 | -78.7921
2.7012 | -78.7943
2.7016 | -78.7956
2.7019 | -78.7976
2.7023 | -78.7992
2.7027 | -78.8
2.703 | -78.7999
2.7034 | -78.8025
2.7037 | -78.8049
2.7041 | -78.8068
2.7045 | -78.8092
2.7048 | -78.8118
2.7052 | -78.813
2.7055 | -78.814
2.7059 | -78.8176
2.7063 | -78.8203
2.7066 | -78.8237
2.707 | -78.8278
2.7074 | -78.83
2.7077 | -78.8324
2.7081 | -78.8349
2.7084 | -78.8371
2.7088 | -78.8395
2.7092 | -78.8418
2.7095 | -78.8443
2.7099 | -78.847
2.7103 | -78.8503
2.7106 | -78.8526
2.711 | -78.8557
2.7113 | -78.8591
2.7117 | -78.8626
2.7121 | -78.8632
2.7124 | -78.8646
2.7128 | -78.8663
2.7131 | -78.8683
2.7135 | -78.8709
2.7139 | -78.8739
2.7142 | -78.8765
2.7146 | -78.879
2.715 | -78.8806
2.7153 | -78.8828
2.7157 | -78.8855
2.716 | -78.8865
2.7164 | -78.8883
2.7168 | -78.8913
2.7171 | -78.8942
2.7175 | -78.8953
2.7178 | -78.897
2.7182 | -78.8993
2.7186 | -78.9012
2.7189 | -78.9036
2.7193 | -78.9058
2.7197 | -78.9079
2.72 | -78.9106
2.7204 | -78.9137
2.7207 | -78.9169
2.7211 | -78.9202
2.7215 | -78.9234
2.7218 | -78.9267
2.7222 | -78.9305
2.7226 | -78.9329
2.7229 | -78.937
2.7233 | -78.9409
2.7236 | -78.9413
2.724 | -78.944
2.7244 | -78.9466
2.7247 | -78.9515
2.7251 | -78.9565
2.7254 | -78.9602
2.7258 | -78.9646
2.7262 | -78.9653
2.7265 | -78.9623
2.7269 | -78.9617
2.7273 | -78.9635
2.7276 | -78.9661
2.728 | -78.9686
2.7283 | -78.9702
2.7287 | -78.9726
2.7291 | -78.9745
2.7294 | -78.9778
2.7298 | -78.9829
2.7302 | -78.9874
2.7305 | -78.9917
2.7309 | -78.9946
2.7312 | -78.9952
2.7316 | -78.9976
2.732 | -78.9985
2.7323 | -78.9973
2.7327 | -78.9984
2.733 | -78.9998
2.7334 | -78.9991
2.7338 | -78.9971
2.7341 | -78.9977
2.7345 | -79.0001
2.7349 | -79.0024
2.7352 | -79.0049
2.7356 | -79.0072
2.7359 | -79.01
2.7363 | -79.0123
2.7367 | -79.0155
2.737 | -79.0221
2.7374 | -79.0268
2.7377 | -79.0264
2.7381 | -79.0254
2.7385 | -79.0245
2.7388 | -79.0254
2.7392 | -79.0267
2.7396 | -79.0289
2.7399 | -79.0289
2.7403 | -79.028
2.7406 | -79.0282
2.741 | -79.0279
2.7414 | -79.0296
2.7417 | -79.0317
2.7421 | -79.033
2.7425 | -79.0336
2.7428 | -79.0336
2.7432 | -79.0331
2.7435 | -79.0328
2.7439 | -79.0345
2.7443 | -79.0362
2.7446 | -79.0353
2.745 | -79.0325
2.7453 | -79.0304
2.7457 | -79.0315
2.7461 | -79.0332
2.7464 | -79.0352
2.7468 | -79.0368
2.7472 | -79.0379
2.7475 | -79.0396
2.7479 | -79.042
2.7482 | -79.0439
2.7486 | -79.0441
2.749 | -79.0482
2.7493 | -79.052
2.7497 | -79.0543
2.75 | -79.0557
2.7504 | -79.057
2.7508 | -79.06
2.7511 | -79.0624
2.7515 | -79.064
2.7519 | -79.0656
2.7522 | -79.0674
2.7526 | -79.0697
2.7529 | -79.072
2.7533 | -79.0738
2.7537 | -79.0743
2.754 | -79.0732
2.7544 | -79.0714
2.7548 | -79.0718
2.7551 | -79.0721
2.7555 | -79.0696
2.7558 | -79.0669
2.7562 | -79.066
2.7566 | -79.0652
2.7569 | -79.0655
2.7573 | -79.0667
2.7576 | -79.0665
2.758 | -79.0664
2.7584 | -79.0665
2.7587 | -79.0668
2.7591 | -79.0673
2.7595 | -79.0679
2.7598 | -79.0685
2.7602 | -79.0674
2.7605 | -79.0661
2.7609 | -79.0657
2.7613 | -79.0678
2.7616 | -79.0683
2.762 | -79.0676
2.7624 | -79.0663
2.7627 | -79.0683
2.7631 | -79.0673
2.7634 | -79.0665
2.7638 | -79.0647
2.7642 | -79.0628
2.7645 | -79.0613
2.7649 | -79.0594
2.7652 | -79.0571
2.7656 | -79.0564
2.766 | -79.055
2.7663 | -79.055
2.7667 | -79.0541
2.7671 | -79.0523
2.7674 | -79.0499
2.7678 | -79.0482
2.7681 | -79.046
2.7685 | -79.0445
2.7689 | -79.0429
2.7692 | -79.041
2.7696 | -79.0395
2.7699 | -79.0376
2.7703 | -79.0335
2.7707 | -79.0294
2.771 | -79.0264
2.7714 | -79.0244
2.7718 | -79.0232
2.7721 | -79.0215
2.7725 | -79.0216
2.7728 | -79.0228
2.7732 | -79.0233
2.7736 | -79.0224
2.7739 | -79.0213
2.7743 | -79.0212
2.7747 | -79.0236
2.775 | -79.0248
2.7754 | -79.0249
2.7757 | -79.0232
2.7761 | -79.0213
2.7765 | -79.0192
2.7768 | -79.019
2.7772 | -79.0195
2.7775 | -79.0173
2.7779 | -79.0159
2.7783 | -79.0169
2.7786 | -79.0165
2.779 | -79.0147
2.7794 | -79.0123
2.7797 | -79.0104
2.7801 | -79.0079
2.7804 | -79.0071
2.7808 | -79.0051
2.7812 | -79.0014
2.7815 | -78.9986
2.7819 | -78.9954
2.7822 | -78.993
2.7826 | -78.9902
2.783 | -78.987
2.7833 | -78.9847
2.7837 | -78.982
2.7841 | -78.9803
2.7844 | -78.9786
2.7848 | -78.9762
2.7851 | -78.9743
2.7855 | -78.9733
2.7859 | -78.9725
2.7862 | -78.971
2.7866 | -78.9688
2.787 | -78.9668
2.7873 | -78.9636
2.7877 | -78.9606
2.788 | -78.9574
2.7884 | -78.9537
2.7888 | -78.9512
2.7891 | -78.9488
2.7895 | -78.9474
2.7898 | -78.946
2.7902 | -78.9442
2.7906 | -78.943
2.7909 | -78.9409
2.7913 | -78.9378
2.7917 | -78.9359
2.792 | -78.9341
2.7924 | -78.9321
2.7927 | -78.9315
2.7931 | -78.9323
2.7935 | -78.9302
2.7938 | -78.9279
2.7942 | -78.925
2.7946 | -78.9221
2.7949 | -78.9202
2.7953 | -78.9184
2.7956 | -78.9163
2.796 | -78.9148
2.7964 | -78.9136
2.7967 | -78.9115
2.7971 | -78.9094
2.7974 | -78.9087
2.7978 | -78.9068
2.7982 | -78.9055
2.7985 | -78.9039
2.7989 | -78.9025
2.7993 | -78.9006
2.7996 | -78.8986
2.8 | -78.897
2.8003 | -78.895
2.8007 | -78.8934
2.8011 | -78.8918
2.8014 | -78.8897
2.8018 | -78.8861
2.8021 | -78.8828
2.8025 | -78.8813
2.8029 | -78.8795
2.8032 | -78.8787
2.8036 | -78.8771
2.804 | -78.8756
2.8043 | -78.8733
2.8047 | -78.871
2.805 | -78.8683
2.8054 | -78.8665
2.8058 | -78.8653
2.8061 | -78.8633
2.8065 | -78.8597
2.8069 | -78.8562
2.8072 | -78.8568
2.8076 | -78.859
2.8079 | -78.8591
2.8083 | -78.8581
2.8087 | -78.8556
2.809 | -78.8533
2.8094 | -78.8499
2.8097 | -78.8487
2.8101 | -78.8495
2.8105 | -78.8504
2.8108 | -78.8509
2.8112 | -78.8509
2.8116 | -78.8488
2.8119 | -78.8463
2.8123 | -78.8427
2.8126 | -78.8386
2.813 | -78.8355
2.8134 | -78.8334
2.8137 | -78.8314
2.8141 | -78.8286
2.8144 | -78.8265
2.8148 | -78.8234
2.8152 | -78.8213
2.8155 | -78.8201
2.8159 | -78.8195
2.8163 | -78.8194
2.8166 | -78.8198
2.817 | -78.8206
2.8173 | -78.8221
2.8177 | -78.8216
2.8181 | -78.8199
2.8184 | -78.82
2.8188 | -78.8186
2.8192 | -78.8171
2.8195 | -78.8155
2.8199 | -78.8129
2.8202 | -78.8116
2.8206 | -78.8107
2.821 | -78.811
2.8213 | -78.8127
2.8217 | -78.814
2.822 | -78.8126
2.8224 | -78.8121
2.8228 | -78.8108
2.8231 | -78.808
2.8235 | -78.8035
2.8239 | -78.8003
2.8242 | -78.799
2.8246 | -78.797
2.8249 | -78.7963
2.8253 | -78.7958
2.8257 | -78.7958
2.826 | -78.7969
2.8264 | -78.797
2.8268 | -78.7956
2.8271 | -78.7949
2.8275 | -78.795
2.8278 | -78.7975
2.8282 | -78.7994
2.8286 | -78.7991
2.8289 | -78.8012
2.8293 | -78.8007
2.8296 | -78.8004
2.83 | -78.8002
2.8304 | -78.7997
2.8307 | -78.7996
2.8311 | -78.7998
2.8315 | -78.7994
2.8318 | -78.7975
2.8322 | -78.7961
2.8325 | -78.7968
2.8329 | -78.7958
2.8333 | -78.7969
2.8336 | -78.7972
2.834 | -78.7978
2.8343 | -78.7983
2.8347 | -78.7989
2.8351 | -78.7992
2.8354 | -78.802
2.8358 | -78.8059
2.8362 | -78.8098
2.8365 | -78.8133
2.8369 | -78.8147
2.8372 | -78.8151
2.8376 | -78.8155
2.838 | -78.8163
2.8383 | -78.8157
2.8387 | -78.816
2.8391 | -78.8186
2.8394 | -78.821
2.8398 | -78.8242
2.8401 | -78.8276
2.8405 | -78.8296
2.8409 | -78.831
2.8412 | -78.8319
2.8416 | -78.8324
2.8419 | -78.8339
2.8423 | -78.8347
2.8427 | -78.8331
2.843 | -78.8341
2.8434 | -78.8352
2.8438 | -78.8365
2.8441 | -78.8372
2.8445 | -78.8382
2.8448 | -78.8396
2.8452 | -78.841
2.8456 | -78.8424
2.8459 | -78.8445
2.8463 | -78.8452
2.8467 | -78.847
2.847 | -78.8472
2.8474 | -78.8482
2.8477 | -78.85
2.8481 | -78.8515
2.8485 | -78.8524
2.8488 | -78.8537
2.8492 | -78.8564
2.8495 | -78.8599
2.8499 | -78.8628
2.8503 | -78.863
2.8506 | -78.863
2.851 | -78.8642
2.8514 | -78.865
2.8517 | -78.8666
2.8521 | -78.8688
2.8524 | -78.8703
2.8528 | -78.8716
2.8532 | -78.8732
2.8535 | -78.8751
2.8539 | -78.8767
2.8542 | -78.8787
2.8546 | -78.8811
2.855 | -78.8849
2.8553 | -78.8886
2.8557 | -78.8895
2.8561 | -78.8898
2.8564 | -78.8905
2.8568 | -78.8934
2.8571 | -78.8953
2.8575 | -78.897
2.8579 | -78.8969
2.8582 | -78.8976
2.8586 | -78.9001
2.859 | -78.9023
2.8593 | -78.9033
2.8597 | -78.905
2.86 | -78.9079
2.8604 | -78.9084
2.8608 | -78.9098
2.8611 | -78.9115
2.8615 | -78.9141
2.8618 | -78.9162
2.8622 | -78.9163
2.8626 | -78.9168
2.8629 | -78.9199
2.8633 | -78.9221
2.8637 | -78.9245
2.864 | -78.9278
2.8644 | -78.9299
2.8647 | -78.9287
2.8651 | -78.929
2.8655 | -78.9308
2.8658 | -78.9325
2.8662 | -78.9348
2.8666 | -78.937
2.8669 | -78.9383
2.8673 | -78.9397
2.8676 | -78.9424
2.868 | -78.9443
2.8684 | -78.9448
2.8687 | -78.9465
2.8691 | -78.9483
2.8694 | -78.951
2.8698 | -78.9541
2.8702 | -78.9561
2.8705 | -78.9585
2.8709 | -78.9597
2.8713 | -78.9602
2.8716 | -78.9602
2.872 | -78.9602
2.8723 | -78.9596
2.8727 | -78.9584
2.8731 | -78.9587
2.8734 | -78.9576
2.8738 | -78.9582
2.8741 | -78.9609
2.8745 | -78.9635
2.8749 | -78.9634
2.8752 | -78.9658
2.8756 | -78.9696
2.876 | -78.9729
2.8763 | -78.9752
2.8767 | -78.9768
2.877 | -78.9783
2.8774 | -78.9807
2.8778 | -78.9816
2.8781 | -78.9834
2.8785 | -78.9853
2.8789 | -78.9862
2.8792 | -78.9871
2.8796 | -78.9876
2.8799 | -78.9885
2.8803 | -78.9886
2.8807 | -78.9895
2.881 | -78.9904
2.8814 | -78.9927
2.8817 | -78.9967
2.8821 | -78.9929
2.8825 | -78.9891
2.8828 | -78.9883
2.8832 | -78.9893
2.8836 | -78.9916
2.8839 | -78.9927
2.8843 | -78.9929
2.8846 | -78.9932
2.885 | -78.9928
2.8854 | -78.9935
2.8857 | -78.9928
2.8861 | -78.9922
2.8864 | -78.9876
2.8868 | -78.9872
2.8872 | -78.9888
2.8875 | -78.9908
2.8879 | -78.9916
2.8883 | -78.993
2.8886 | -78.9908
2.889 | -78.9901
2.8893 | -78.9906
2.8897 | -78.9901
2.8901 | -78.99
2.8904 | -78.99
2.8908 | -78.9895
2.8912 | -78.9894
2.8915 | -78.9888
2.8919 | -78.9868
2.8922 | -78.985
2.8926 | -78.9848
2.893 | -78.9838
2.8933 | -78.9835
2.8937 | -78.9833
2.894 | -78.982
2.8944 | -78.977
2.8948 | -78.9728
2.8951 | -78.9712
2.8955 | -78.9706
2.8959 | -78.9689
2.8962 | -78.9672
2.8966 | -78.9663
2.8969 | -78.965
2.8973 | -78.9636
2.8977 | -78.9646
2.898 | -78.966
2.8984 | -78.9655
2.8988 | -78.9649
2.8991 | -78.9637
2.8995 | -78.9607
2.8998 | -78.9588
2.9002 | -78.9583
2.9006 | -78.9581
2.9009 | -78.9564
2.9013 | -78.9558
2.9016 | -78.9536
2.902 | -78.9501
2.9024 | -78.9482
2.9027 | -78.9474
2.9031 | -78.9441
2.9035 | -78.939
2.9038 | -78.9366
2.9042 | -78.935
2.9045 | -78.9338
2.9049 | -78.9327
2.9053 | -78.9306
2.9056 | -78.9295
2.906 | -78.9286
2.9063 | -78.9277
2.9067 | -78.9273
2.9071 | -78.9266
2.9074 | -78.9263
2.9078 | -78.9255
2.9082 | -78.9249
2.9085 | -78.9236
2.9089 | -78.921
2.9092 | -78.9198
2.9096 | -78.9188
2.91 | -78.9171
2.9103 | -78.916
2.9107 | -78.9151
2.9111 | -78.913
2.9114 | -78.9117
2.9118 | -78.9118
2.9121 | -78.9109
2.9125 | -78.9105
2.9129 | -78.9083
2.9132 | -78.9057
2.9136 | -78.9055
2.9139 | -78.9026
2.9143 | -78.9006
2.9147 | -78.8997
2.915 | -78.898
2.9154 | -78.8965
2.9158 | -78.8957
2.9161 | -78.8943
2.9165 | -78.8911
2.9168 | -78.8889
2.9172 | -78.8852
2.9176 | -78.881
2.9179 | -78.8786
2.9183 | -78.8766
2.9186 | -78.8745
2.919 | -78.8734
2.9194 | -78.8721
2.9197 | -78.8729
2.9201 | -78.8713
2.9205 | -78.8683
2.9208 | -78.8646
2.9212 | -78.8612
2.9215 | -78.8589
2.9219 | -78.8569
2.9223 | -78.8538
2.9226 | -78.8511
2.923 | -78.8507
2.9234 | -78.8521
2.9237 | -78.8522
2.9241 | -78.8514
2.9244 | -78.8499
2.9248 | -78.8484
2.9252 | -78.8467
2.9255 | -78.8446
2.9259 | -78.843
2.9262 | -78.8415
2.9266 | -78.8385
2.927 | -78.8352
2.9273 | -78.832
2.9277 | -78.8298
2.9281 | -78.8297
2.9284 | -78.828
2.9288 | -78.8276
2.9291 | -78.8288
2.9295 | -78.8275
2.9299 | -78.8257
2.9302 | -78.8259
2.9306 | -78.8262
2.931 | -78.8228
2.9313 | -78.8203
2.9317 | -78.8179
2.932 | -78.8153
2.9324 | -78.8123
2.9328 | -78.8096
2.9331 | -78.8085
2.9335 | -78.8051
2.9338 | -78.8025
2.9342 | -78.8003
2.9346 | -78.7978
2.9349 | -78.7962
2.9353 | -78.7962
2.9357 | -78.796
2.936 | -78.7948
2.9364 | -78.7947
2.9367 | -78.7935
2.9371 | -78.7915
2.9375 | -78.7909
2.9378 | -78.791
2.9382 | -78.7899
2.9385 | -78.7893
2.9389 | -78.7881
2.9393 | -78.786
2.9396 | -78.784
2.94 | -78.7829
2.9404 | -78.7811
2.9407 | -78.7801
2.9411 | -78.7797
2.9414 | -78.7764
2.9418 | -78.7717
2.9422 | -78.7706
2.9425 | -78.7712
2.9429 | -78.7709
2.9433 | -78.769
2.9436 | -78.7672
2.944 | -78.7642
2.9443 | -78.7599
2.9447 | -78.7574
2.9451 | -78.7548
2.9454 | -78.7512
2.9458 | -78.7473
2.9461 | -78.7446
2.9465 | -78.742
2.9469 | -78.7399
2.9472 | -78.7376
2.9476 | -78.7355
2.948 | -78.7331
2.9483 | -78.7322
2.9487 | -78.7306
2.949 | -78.7283
2.9494 | -78.7244
2.9498 | -78.7222
2.9501 | -78.7203
2.9505 | -78.7179
2.9508 | -78.716
2.9512 | -78.7152
2.9516 | -78.7158
2.9519 | -78.716
2.9523 | -78.7156
2.9527 | -78.7166
2.953 | -78.7161
2.9534 | -78.715
2.9537 | -78.7157
2.9541 | -78.7145
2.9545 | -78.7128
2.9548 | -78.7112
2.9552 | -78.7091
2.9556 | -78.7065
2.9559 | -78.7041
2.9563 | -78.702
2.9566 | -78.7012
2.957 | -78.6999
2.9574 | -78.6973
2.9577 | -78.6952
2.9581 | -78.6944
2.9584 | -78.6931
2.9588 | -78.6914
2.9592 | -78.691
2.9595 | -78.6909
2.9599 | -78.6898
2.9603 | -78.6883
2.9606 | -78.6856
2.961 | -78.6811
2.9613 | -78.6802
2.9617 | -78.6797
2.9621 | -78.6799
2.9624 | -78.6789
2.9628 | -78.6774
2.9632 | -78.6765
2.9635 | -78.6759
2.9639 | -78.6761
2.9642 | -78.6759
2.9646 | -78.678
2.965 | -78.6769
2.9653 | -78.6758
2.9657 | -78.6741
2.966 | -78.671
2.9664 | -78.6692
2.9668 | -78.669
2.9671 | -78.6687
2.9675 | -78.6672
2.9679 | -78.6662
2.9682 | -78.6633
2.9686 | -78.6628
2.9689 | -78.6614
2.9693 | -78.6603
2.9697 | -78.6602
2.97 | -78.6623
2.9704 | -78.6643
2.9707 | -78.6652
2.9711 | -78.6658
2.9715 | -78.6653
2.9718 | -78.6653
2.9722 | -78.6653
2.9726 | -78.665
2.9729 | -78.6654
2.9733 | -78.6672
2.9736 | -78.6684
2.974 | -78.6697
2.9744 | -78.6702
2.9747 | -78.6716
2.9751 | -78.6719
2.9755 | -78.6709
2.9758 | -78.6699
2.9762 | -78.67
2.9765 | -78.6699
2.9769 | -78.6703
2.9773 | -78.6706
2.9776 | -78.6701
2.978 | -78.67
2.9783 | -78.6706
2.9787 | -78.6709
2.9791 | -78.6714
2.9794 | -78.6718
2.9798 | -78.6722
2.9802 | -78.6729
2.9805 | -78.6738
2.9809 | -78.6745
2.9812 | -78.677
2.9816 | -78.6788
2.982 | -78.6787
2.9823 | -78.6797
2.9827 | -78.6803
2.983 | -78.6811
2.9834 | -78.6829
2.9838 | -78.6842
2.9841 | -78.6845
2.9845 | -78.6841
2.9849 | -78.6852
2.9852 | -78.6898
2.9856 | -78.6913
2.9859 | -78.6928
2.9863 | -78.6941
2.9867 | -78.6956
2.987 | -78.697
2.9874 | -78.6987
2.9878 | -78.7
2.9881 | -78.7022
2.9885 | -78.7048
2.9888 | -78.7066
2.9892 | -78.7084
2.9896 | -78.7104
2.9899 | -78.7119
2.9903 | -78.7139
2.9906 | -78.7154
2.991 | -78.7174
2.9914 | -78.7185
2.9917 | -78.7206
2.9921 | -78.7231
2.9925 | -78.7256
2.9928 | -78.7283
2.9932 | -78.7311
2.9935 | -78.7341
2.9939 | -78.7372
2.9943 | -78.7399
2.9946 | -78.7418
2.995 | -78.7439
2.9954 | -78.7471
2.9957 | -78.7524
2.9961 | -78.7543
2.9964 | -78.7568
2.9968 | -78.7607
2.9972 | -78.7644
2.9975 | -78.7675
2.9979 | -78.7691
2.9982 | -78.7715
2.9986 | -78.7748
2.999 | -78.7782
2.9993 | -78.7812
2.9997 | -78.7844
3.0001 | -78.7869
3.0004 | -78.7896
3.0008 | -78.7925
3.0011 | -78.7956
3.0015 | -78.7994
3.0019 | -78.8027
3.0022 | -78.806
3.0026 | -78.808
3.0029 | -78.8096
3.0033 | -78.8143
3.0037 | -78.8205
3.004 | -78.8241
3.0044 | -78.8252
3.0048 | -78.8262
3.0051 | -78.8277
3.0055 | -78.8317
3.0058 | -78.8358
3.0062 | -78.8418
3.0066 | -78.8458
3.0069 | -78.8497
3.0073 | -78.854
3.0077 | -78.8552
3.008 | -78.8578
3.0084 | -78.8633
3.0087 | -78.8702
3.0091 | -78.8771
3.0095 | -78.8808
3.0098 | -78.8837
3.0102 | -78.8874
3.0105 | -78.8907
3.0109 | -78.8935
3.0113 | -78.8972
3.0116 | -78.9012
3.012 | -78.9058
3.0124 | -78.9106
3.0127 | -78.9139
3.0131 | -78.9163
3.0134 | -78.9193
3.0138 | -78.923
3.0142 | -78.9284
3.0145 | -78.9336
3.0149 | -78.9399
3.0152 | -78.9469
3.0156 | -78.9531
3.016 | -78.959
3.0163 | -78.9645
3.0167 | -78.9707
3.0171 | -78.9735
3.0174 | -78.9765
3.0178 | -78.9825
3.0181 | -78.9879
3.0185 | -78.992
3.0189 | -78.9958
3.0192 | -79
3.0196 | -79.0042
3.02 | -79.0075
3.0203 | -79.0119
3.0207 | -79.0165
3.021 | -79.0215
3.0214 | -79.028
3.0218 | -79.0328
3.0221 | -79.0375
3.0225 | -79.0429
3.0228 | -79.0482
3.0232 | -79.055
3.0236 | -79.0637
3.0239 | -79.0707
3.0243 | -79.0751
3.0247 | -79.0788
3.025 | -79.0827
3.0254 | -79.0895
3.0257 | -79.095
3.0261 | -79.0992
3.0265 | -79.1042
3.0268 | -79.1123
3.0272 | -79.118
3.0276 | -79.1218
3.0279 | -79.1267
3.0283 | -79.1323
3.0286 | -79.1362
3.029 | -79.1397
3.0294 | -79.1432
3.0297 | -79.1473
3.0301 | -79.1493
3.0304 | -79.1523
3.0308 | -79.1564
3.0312 | -79.1615
3.0315 | -79.1659
3.0319 | -79.1711
3.0323 | -79.1751
3.0326 | -79.1789
3.033 | -79.183
3.0333 | -79.1869
3.0337 | -79.1898
3.0341 | -79.1942
3.0344 | -79.1994
3.0348 | -79.2041
3.0351 | -79.2084
3.0355 | -79.2136
3.0359 | -79.2201
3.0362 | -79.2261
3.0366 | -79.2299
3.037 | -79.2353
3.0373 | -79.2396
3.0377 | -79.2434
3.038 | -79.248
3.0384 | -79.2518
3.0388 | -79.2558
3.0391 | -79.2598
3.0395 | -79.2646
3.0399 | -79.2691
3.0402 | -79.2747
3.0406 | -79.2803
3.0409 | -79.2836
3.0413 | -79.287
3.0417 | -79.2918
3.042 | -79.297
3.0424 | -79.3024
3.0427 | -79.3084
3.0431 | -79.3141
3.0435 | -79.3195
3.0438 | -79.3241
3.0442 | -79.3285
3.0446 | -79.3327
3.0449 | -79.3372
3.0453 | -79.3413
3.0456 | -79.3456
3.046 | -79.3501
3.0464 | -79.3548
3.0467 | -79.3594
3.0471 | -79.365
3.0474 | -79.3722
3.0478 | -79.3773
3.0482 | -79.3819
3.0485 | -79.3866
3.0489 | -79.3911
3.0493 | -79.3952
3.0496 | -79.4001
3.05 | -79.4041
3.0503 | -79.4084
3.0507 | -79.4126
3.0511 | -79.4156
3.0514 | -79.4206
3.0518 | -79.4241
3.0522 | -79.4283
3.0525 | -79.4328
3.0529 | -79.4364
3.0532 | -79.4389
3.0536 | -79.4419
3.054 | -79.4458
3.0543 | -79.4507
3.0547 | -79.4552
3.055 | -79.4595
3.0554 | -79.4633
3.0558 | -79.469
3.0561 | -79.4753
3.0565 | -79.4794
3.0569 | -79.4838
3.0572 | -79.4882
3.0576 | -79.4922
3.0579 | -79.4962
3.0583 | -79.4968
3.0587 | -79.5011
3.059 | -79.5059
3.0594 | -79.5095
3.0598 | -79.5103
3.0601 | -79.5122
3.0605 | -79.5157
3.0608 | -79.5196
3.0612 | -79.5235
3.0616 | -79.5279
3.0619 | -79.5334
3.0623 | -79.5397
3.0626 | -79.5445
3.063 | -79.5484
3.0634 | -79.553
3.0637 | -79.5594
3.0641 | -79.5638
3.0645 | -79.567
3.0648 | -79.5704
3.0652 | -79.5745
3.0655 | -79.5764
3.0659 | -79.5766
3.0663 | -79.5763
3.0666 | -79.5794
3.067 | -79.5833
3.0673 | -79.5864
3.0677 | -79.5907
3.0681 | -79.5955
3.0684 | -79.6006
3.0688 | -79.6032
3.0692 | -79.6076
3.0695 | -79.6101
3.0699 | -79.613
3.0702 | -79.6157
3.0706 | -79.6185
3.071 | -79.6223
3.0713 | -79.629
3.0717 | -79.6321
3.0721 | -79.6351
3.0724 | -79.6374
3.0728 | -79.6407
3.0731 | -79.6439
3.0735 | -79.6478
3.0739 | -79.6509
3.0742 | -79.6538
3.0746 | -79.6581
3.0749 | -79.6626
3.0753 | -79.6661
3.0757 | -79.6689
3.076 | -79.6718
3.0764 | -79.6746
3.0768 | -79.6773
3.0771 | -79.6802
3.0775 | -79.6828
3.0778 | -79.6859
3.0782 | -79.6895
3.0786 | -79.6943
3.0789 | -79.6987
3.0793 | -79.7002
3.0797 | -79.705
3.08 | -79.7131
3.0804 | -79.7201
3.0807 | -79.7254
3.0811 | -79.7279
3.0815 | -79.7294
3.0818 | -79.7304
3.0822 | -79.7337
3.0825 | -79.7364
3.0829 | -79.7385
3.0833 | -79.7405
3.0836 | -79.7424
3.084 | -79.7442
3.0844 | -79.7456
3.0847 | -79.7468
3.0851 | -79.7493
3.0854 | -79.7519
3.0858 | -79.7526
3.0862 | -79.7544
3.0865 | -79.7567
3.0869 | -79.7592
3.0872 | -79.7623
3.0876 | -79.7657
3.088 | -79.7691
3.0883 | -79.7718
3.0887 | -79.7739
3.0891 | -79.7766
3.0894 | -79.7793
3.0898 | -79.7815
3.0901 | -79.7912
3.0905 | -79.7998
3.0909 | -79.8036
3.0912 | -79.8061
3.0916 | -79.8076
3.092 | -79.8091
3.0923 | -79.8113
3.0927 | -79.8127
3.093 | -79.814
3.0934 | -79.8133
3.0938 | -79.8146
3.0941 | -79.818
3.0945 | -79.8258
3.0948 | -79.8296
3.0952 | -79.8311
3.0956 | -79.8311
3.0959 | -79.8311
3.0963 | -79.8323
3.0967 | -79.8359
3.097 | -79.8399
3.0974 | -79.8423
3.0977 | -79.8437
3.0981 | -79.8443
3.0985 | -79.8433
3.0988 | -79.8417
3.0992 | -79.8414
3.0995 | -79.8418
3.0999 | -79.8421
3.1003 | -79.8422
3.1006 | -79.8466
3.101 | -79.8473
3.1014 | -79.8474
3.1017 | -79.849
3.1021 | -79.8512
3.1024 | -79.8526
3.1028 | -79.8548
3.1032 | -79.8621
3.1035 | -79.8684
3.1039 | -79.8711
3.1043 | -79.8727
3.1046 | -79.8742
3.105 | -79.8771
3.1053 | -79.8787
3.1057 | -79.8807
3.1061 | -79.8844
3.1064 | -79.8863
3.1068 | -79.8875
3.1071 | -79.8884
3.1075 | -79.8882
3.1079 | -79.8902
3.1082 | -79.8919
3.1086 | -79.8926
3.109 | -79.8946
3.1093 | -79.8966
3.1097 | -79.8961
3.11 | -79.896
3.1104 | -79.8969
3.1108 | -79.8981
3.1111 | -79.9011
3.1115 | -79.9048
3.1119 | -79.9064
3.1122 | -79.9085
3.1126 | -79.9124
3.1129 | -79.9168
3.1133 | -79.9165
3.1137 | -79.9158
3.114 | -79.9168
3.1144 | -79.9181
3.1147 | -79.9201
3.1151 | -79.9208
3.1155 | -79.9221
3.1158 | -79.9227
3.1162 | -79.9234
3.1166 | -79.9235
3.1169 | -79.924
3.1173 | -79.9248
3.1176 | -79.9247
3.118 | -79.9259
3.1184 | -79.9282
3.1187 | -79.9298
3.1191 | -79.9312
3.1194 | -79.933
3.1198 | -79.9337
3.1202 | -79.9336
3.1205 | -79.9346
3.1209 | -79.9343
3.1213 | -79.934
3.1216 | -79.9335
3.122 | -79.9328
3.1223 | -79.933
3.1227 | -79.9342
3.1231 | -79.9361
3.1234 | -79.9387
3.1238 | -79.9391
3.1242 | -79.9432
3.1245 | -79.944
3.1249 | -79.9446
3.1252 | -79.9456
3.1256 | -79.9437
3.126 | -79.9411
3.1263 | -79.9404
3.1267 | -79.9428
3.127 | -79.9443
3.1274 | -79.9465
3.1278 | -79.9489
3.1281 | -79.95
3.1285 | -79.9505
3.1289 | -79.953
3.1292 | -79.954
3.1296 | -79.9536
3.1299 | -79.9525
3.1303 | -79.9521
3.1307 | -79.9533
3.131 | -79.9569
3.1314 | -79.9575
3.1317 | -79.9608
3.1321 | -79.9607
3.1325 | -79.9606
3.1328 | -79.9616
3.1332 | -79.9613
3.1336 | -79.962
3.1339 | -79.9628
3.1343 | -79.9652
3.1346 | -79.9659
3.135 | -79.9656
3.1354 | -79.965
3.1357 | -79.9655
3.1361 | -79.9662
3.1365 | -79.9669
3.1368 | -79.9678
3.1372 | -79.9674
3.1375 | -79.9674
3.1379 | -79.9686
3.1383 | -79.9697
3.1386 | -79.9684
3.139 | -79.9676
3.1393 | -79.9657
3.1397 | -79.9636
3.1401 | -79.964
3.1404 | -79.9622
3.1408 | -79.9592
3.1412 | -79.9565
3.1415 | -79.956
3.1419 | -79.9553
3.1422 | -79.953
3.1426 | -79.9483
3.143 | -79.9441
3.1433 | -79.9417
3.1437 | -79.9411
3.1441 | -79.937
3.1444 | -79.9299
3.1448 | -79.9248
3.1451 | -79.915
3.1455 | -79.9056
3.1459 | -79.9038
3.1462 | -79.9007
3.1466 | -79.8967
3.1469 | -79.8939
3.1473 | -79.8884
3.1477 | -79.8828
3.148 | -79.8805
3.1484 | -79.8787
3.1488 | -79.8745
3.1491 | -79.8674
3.1495 | -79.8651
3.1498 | -79.8633
3.1502 | -79.8588
3.1506 | -79.8546
3.1509 | -79.85
3.1513 | -79.8466
3.1516 | -79.8446
3.152 | -79.8427
3.1524 | -79.8367
3.1527 | -79.8303
3.1531 | -79.8263
3.1535 | -79.8243
3.1538 | -79.8192
3.1542 | -79.8081
3.1545 | -79.8045
3.1549 | -79.7995
3.1553 | -79.7943
3.1556 | -79.7885
3.156 | -79.7824
3.1564 | -79.7736
3.1567 | -79.7712
3.1571 | -79.7652
3.1574 | -79.7594
3.1578 | -79.7525
3.1582 | -79.7471
3.1585 | -79.7367
3.1589 | -79.7316
3.1592 | -79.7277
3.1596 | -79.7202
3.16 | -79.7131
3.1603 | -79.7071
3.1607 | -79.6964
3.1611 | -79.6857
3.1614 | -79.6783
3.1618 | -79.6734
3.1621 | -79.6673
3.1625 | -79.6618
3.1629 | -79.6385
3.1632 | -79.6276
3.1636 | -79.6217
3.1639 | -79.6139
3.1643 | -79.608
3.1647 | -79.5994
3.165 | -79.5859
3.1654 | -79.578
3.1658 | -79.5706
3.1661 | -79.5653
3.1665 | -79.5585
3.1668 | -79.5519
3.1672 | -79.5467
3.1676 | -79.5406
3.1679 | -79.534
3.1683 | -79.5278
3.1687 | -79.5216
3.169 | -79.5125
3.1694 | -79.5038
3.1697 | -79.4963
3.1701 | -79.489
3.1705 | -79.4808
3.1708 | -79.4736
3.1712 | -79.4657
3.1715 | -79.4565
3.1719 | -79.448
3.1723 | -79.4397
3.1726 | -79.4313
3.173 | -79.4232
3.1734 | -79.4116
3.1737 | -79.4027
3.1741 | -79.3953
3.1744 | -79.3909
3.1748 | -79.3837
3.1752 | -79.3749
3.1755 | -79.3666
3.1759 | -79.3577
3.1763 | -79.3493
3.1766 | -79.3413
3.177 | -79.3326
3.1773 | -79.3229
3.1777 | -79.3132
3.1781 | -79.304
3.1784 | -79.2944
3.1788 | -79.2858
3.1791 | -79.2773
3.1795 | -79.2694
3.1799 | -79.2623
3.1802 | -79.2539
3.1806 | -79.2448
3.181 | -79.2363
3.1813 | -79.2295
3.1817 | -79.222
3.182 | -79.2148
3.1824 | -79.2047
3.1828 | -79.195
3.1831 | -79.1852
3.1835 | -79.1711
3.1838 | -79.1612
3.1842 | -79.1512
3.1846 | -79.1412
3.1849 | -79.1311
3.1853 | -79.1211
3.1857 | -79.1112
3.186 | -79.0993
3.1864 | -79.0876
3.1867 | -79.0773
3.1871 | -79.0678
3.1875 | -79.0575
3.1878 | -79.0464
3.1882 | -79.0355
3.1886 | -79.0239
3.1889 | -79.014
3.1893 | -79.0041
3.1896 | -78.9936
3.19 | -78.984
3.1904 | -78.9745
3.1907 | -78.9639
3.1911 | -78.9535
3.1914 | -78.9427
3.1918 | -78.9318
3.1922 | -78.9213
3.1925 | -78.9104
3.1929 | -78.8995
3.1933 | -78.8889
3.1936 | -78.8781
3.194 | -78.8646
3.1943 | -78.8538
3.1947 | -78.8435
3.1951 | -78.8328
3.1954 | -78.8221
3.1958 | -78.8112
3.1961 | -78.7999
3.1965 | -78.7892
3.1969 | -78.7794
3.1972 | -78.7685
3.1976 | -78.7572
3.198 | -78.747
3.1983 | -78.7363
3.1987 | -78.7255
3.199 | -78.7161
3.1994 | -78.7069
3.1998 | -78.6962
3.2001 | -78.685
3.2005 | -78.6742
3.2009 | -78.6618
3.2012 | -78.6494
3.2016 | -78.6384
3.2019 | -78.6275
3.2023 | -78.6168
3.2027 | -78.6056
3.203 | -78.595
3.2034 | -78.5843
3.2037 | -78.5733
3.2041 | -78.5624
3.2045 | -78.5502
3.2048 | -78.5402
3.2052 | -78.5304
3.2056 | -78.5199
3.2059 | -78.5092
3.2063 | -78.499
3.2066 | -78.488
3.207 | -78.4775
3.2074 | -78.467
3.2077 | -78.4564
3.2081 | -78.4458
3.2085 | -78.4354
3.2088 | -78.4246
3.2092 | -78.4138
3.2095 | -78.4031
3.2099 | -78.3925
3.2103 | -78.3814
3.2106 | -78.3704
3.211 | -78.3591
3.2113 | -78.3475
3.2117 | -78.3364
3.2121 | -78.3268
3.2124 | -78.3172
3.2128 | -78.307
3.2132 | -78.2963
3.2135 | -78.2859
3.2139 | -78.2765
3.2142 | -78.2658
3.2146 | -78.2555
3.215 | -78.2469
3.2153 | -78.2368
3.2157 | -78.2272
3.216 | -78.2172
3.2164 | -78.2078
3.2168 | -78.1982
3.2171 | -78.1888
3.2175 | -78.1788
3.2179 | -78.1689
3.2182 | -78.159
3.2186 | -78.1498
3.2189 | -78.1413
3.2193 | -78.1337
3.2197 | -78.1243
3.22 | -78.1142
3.2204 | -78.1055
3.2208 | -78.098
3.2211 | -78.0904
3.2215 | -78.0827
3.2218 | -78.0737
3.2222 | -78.0648
3.2226 | -78.0557
3.2229 | -78.0469
3.2233 | -78.0381
3.2236 | -78.029
3.224 | -78.0204
3.2244 | -78.0138
3.2247 | -78.0061
3.2251 | -77.9976
3.2255 | -77.9924
3.2258 | -77.9834
3.2262 | -77.9742
3.2265 | -77.9653
3.2269 | -77.9572
3.2273 | -77.9489
3.2276 | -77.9412
3.228 | -77.9337
3.2284 | -77.9261
3.2287 | -77.9186
3.2291 | -77.9115
3.2294 | -77.9036
3.2298 | -77.8961
3.2302 | -77.8896
3.2305 | -77.8833
3.2309 | -77.8761
3.2312 | -77.8688
3.2316 | -77.8615
3.232 | -77.8544
3.2323 | -77.8468
3.2327 | -77.8393
3.2331 | -77.8324
3.2334 | -77.8258
3.2338 | -77.8194
3.2341 | -77.8131
3.2345 | -77.8088
3.2349 | -77.8036
3.2352 | -77.7963
3.2356 | -77.7895
3.2359 | -77.7876
3.2363 | -77.7819
3.2367 | -77.776
3.237 | -77.7693
3.2374 | -77.7627
3.2378 | -77.7578
3.2381 | -77.7513
3.2385 | -77.7458
3.2388 | -77.7408
3.2392 | -77.7357
3.2396 | -77.7307
3.2399 | -77.7258
3.2403 | -77.7203
3.2407 | -77.7154
3.241 | -77.7109
3.2414 | -77.7062
3.2417 | -77.7016
3.2421 | -77.697
3.2425 | -77.6935
3.2428 | -77.6907
3.2432 | -77.6867
3.2435 | -77.6826
3.2439 | -77.6797
3.2443 | -77.6769
3.2446 | -77.6734
3.245 | -77.6695
3.2454 | -77.6657
3.2457 | -77.663
3.2461 | -77.6598
3.2464 | -77.6632
3.2468 | -77.6603
3.2472 | -77.6578
3.2475 | -77.6553
3.2479 | -77.6532
3.2483 | -77.6508
3.2486 | -77.6502
3.249 | -77.6495
3.2493 | -77.6479
3.2497 | -77.6456
3.2501 | -77.6436
3.2504 | -77.6416
3.2508 | -77.6402
3.2511 | -77.6391
3.2515 | -77.6382
3.2519 | -77.6373
3.2522 | -77.6369
3.2526 | -77.6359
3.253 | -77.6338
3.2533 | -77.6317
3.2537 | -77.6301
3.254 | -77.629
3.2544 | -77.6284
3.2548 | -77.6279
3.2551 | -77.627
3.2555 | -77.6271
3.2558 | -77.627
3.2562 | -77.6272
3.2566 | -77.6331
3.2569 | -77.6334
3.2573 | -77.6331
3.2577 | -77.6333
3.258 | -77.6339
3.2584 | -77.6347
3.2587 | -77.6353
3.2591 | -77.6354
3.2595 | -77.6361
3.2598 | -77.637
3.2602 | -77.6389
3.2606 | -77.6411
3.2609 | -77.6425
3.2613 | -77.6434
3.2616 | -77.6449
3.262 | -77.6466
3.2624 | -77.6476
3.2627 | -77.6487
3.2631 | -77.6506
3.2634 | -77.6525
3.2638 | -77.6551
3.2642 | -77.6577
3.2645 | -77.6602
3.2649 | -77.6624
3.2653 | -77.6652
3.2656 | -77.6683
3.266 | -77.6712
3.2663 | -77.674
3.2667 | -77.6767
3.2671 | -77.6868
3.2674 | -77.6904
3.2678 | -77.6939
3.2681 | -77.6975
3.2685 | -77.7015
3.2689 | -77.7058
3.2692 | -77.7102
3.2696 | -77.7138
3.27 | -77.7183
3.2703 | -77.7229
3.2707 | -77.7278
3.271 | -77.7328
3.2714 | -77.7372
3.2718 | -77.7422
3.2721 | -77.7473
3.2725 | -77.752
3.2729 | -77.7573
3.2732 | -77.7641
3.2736 | -77.7706
3.2739 | -77.776
3.2743 | -77.7815
3.2747 | -77.787
3.275 | -77.7927
3.2754 | -77.7988
3.2757 | -77.8053
3.2761 | -77.8115
3.2765 | -77.8177
3.2768 | -77.8239
3.2772 | -77.8304
3.2776 | -77.8428
3.2779 | -77.8491
3.2783 | -77.8547
3.2786 | -77.8628
3.279 | -77.8705
3.2794 | -77.8779
3.2797 | -77.8849
3.2801 | -77.8921
3.2805 | -77.9
3.2808 | -77.9086
3.2812 | -77.9154
3.2815 | -77.9224
3.2819 | -77.9303
3.2823 | -77.9385
3.2826 | -77.9461
3.283 | -77.9543
3.2833 | -77.9623
3.2837 | -77.9697
3.2841 | -77.9774
3.2844 | -77.985
3.2848 | -77.9931
3.2852 | -78.0017
3.2855 | -78.0099
3.2859 | -78.0178
3.2862 | -78.027
3.2866 | -78.0372
3.287 | -78.0475
3.2873 | -78.0569
3.2877 | -78.0656
3.288 | -78.0792
3.2884 | -78.0885
3.2888 | -78.0977
3.2891 | -78.1071
3.2895 | -78.1158
3.2899 | -78.1248
3.2902 | -78.1345
3.2906 | -78.1443
3.2909 | -78.1541
3.2913 | -78.1639
3.2917 | -78.1736
3.292 | -78.1824
3.2924 | -78.1919
3.2928 | -78.2019
3.2931 | -78.2122
3.2935 | -78.2212
3.2938 | -78.2319
3.2942 | -78.2423
3.2946 | -78.2528
3.2949 | -78.2636
3.2953 | -78.2741
3.2956 | -78.2832
3.296 | -78.2927
3.2964 | -78.3021
3.2967 | -78.3126
3.2971 | -78.3234
3.2975 | -78.3343
3.2978 | -78.3455
3.2982 | -78.3567
3.2985 | -78.3692
3.2989 | -78.38
3.2993 | -78.3911
3.2996 | -78.4028
3.3 | -78.4142
3.3003 | -78.4254
3.3007 | -78.4373
3.3011 | -78.4495
3.3014 | -78.4617
3.3018 | -78.4736
3.3022 | -78.4852
3.3025 | -78.4967
3.3029 | -78.5087
3.3032 | -78.5202
3.3036 | -78.532
3.304 | -78.5447
3.3043 | -78.5571
3.3047 | -78.5694
3.3051 | -78.5815
3.3054 | -78.593
3.3058 | -78.6043
3.3061 | -78.6164
3.3065 | -78.6296
3.3069 | -78.6423
3.3072 | -78.6542
3.3076 | -78.6664
3.3079 | -78.6801
3.3083 | -78.693
3.3087 | -78.7051
3.309 | -78.7171
3.3094 | -78.7297
3.3098 | -78.7422
3.3101 | -78.7547
3.3105 | -78.7672
3.3108 | -78.7798
3.3112 | -78.7925
3.3116 | -78.8046
3.3119 | -78.8173
3.3123 | -78.8298
3.3127 | -78.843
3.313 | -78.8548
3.3134 | -78.8669
3.3137 | -78.8798
3.3141 | -78.8925
3.3145 | -78.9056
3.3148 | -78.9179
3.3152 | -78.931
3.3155 | -78.9439
3.3159 | -78.9581
3.3163 | -78.9736
3.3166 | -78.9879
3.317 | -79.0002
3.3174 | -79.0133
3.3177 | -79.026
3.3181 | -79.039
3.3184 | -79.0532
3.3188 | -79.0673
3.3192 | -79.0801
3.3195 | -79.0915
3.3199 | -79.1045
3.3202 | -79.1169
3.3206 | -79.1298
3.321 | -79.1434
3.3213 | -79.1573
3.3217 | -79.1718
3.3221 | -79.1862
3.3224 | -79.2002
3.3228 | -79.2132
3.3231 | -79.2271
3.3235 | -79.2403
3.3239 | -79.2532
3.3242 | -79.2658
3.3246 | -79.2808
3.325 | -79.2947
3.3253 | -79.3076
3.3257 | -79.3197
3.326 | -79.3314
3.3264 | -79.3432
3.3268 | -79.3554
3.3271 | -79.3677
3.3275 | -79.3796
3.3278 | -79.3912
3.3282 | -79.4043
3.3286 | -79.4162
3.3289 | -79.4287
3.3293 | -79.4415
3.3297 | -79.4496
3.33 | -79.461
3.3304 | -79.4734
3.3307 | -79.4853
3.3311 | -79.4976
3.3315 | -79.5106
3.3318 | -79.5225
3.3322 | -79.5364
3.3325 | -79.5503
3.3329 | -79.5623
3.3333 | -79.5747
3.3336 | -79.589
3.334 | -79.6021
3.3344 | -79.6142
3.3347 | -79.6263
3.3351 | -79.6388
3.3354 | -79.6495
3.3358 | -79.6612
3.3362 | -79.6737
3.3365 | -79.6849
3.3369 | -79.6962
3.3373 | -79.709
3.3376 | -79.723
3.338 | -79.7363
3.3383 | -79.7479
3.3387 | -79.7588
3.3391 | -79.7681
3.3394 | -79.7771
3.3398 | -79.7884
3.3401 | -79.7951
3.3405 | -79.8063
3.3409 | -79.8179
3.3412 | -79.8295
3.3416 | -79.8403
3.342 | -79.8519
3.3423 | -79.8645
3.3427 | -79.8764
3.343 | -79.8867
3.3434 | -79.8959
3.3438 | -79.9056
3.3441 | -79.9154
3.3445 | -79.9259
3.3449 | -79.9377
3.3452 | -79.9495
3.3456 | -79.9616
3.3459 | -79.972
3.3463 | -79.9829
3.3467 | -79.9939
3.347 | -80.0047
3.3474 | -80.0146
3.3477 | -80.0228
3.3481 | -80.0317
3.3485 | -80.0413
3.3488 | -80.0509
3.3492 | -80.0599
3.3496 | -80.0709
3.3499 | -80.0807
3.3503 | -80.0907
3.3506 | -80.0958
3.351 | -80.1069
3.3514 | -80.1182
3.3517 | -80.1286
3.3521 | -80.1371
3.3524 | -80.1452
3.3528 | -80.1538
3.3532 | -80.1628
3.3535 | -80.1714
3.3539 | -80.18
3.3543 | -80.1882
3.3546 | -80.1963
3.355 | -80.206
3.3553 | -80.2162
3.3557 | -80.2269
3.3561 | -80.2362
3.3564 | -80.2446
3.3568 | -80.2537
3.3572 | -80.2617
3.3575 | -80.2692
3.3579 | -80.2777
3.3582 | -80.2864
3.3586 | -80.2942
3.359 | -80.2999
3.3593 | -80.305
3.3597 | -80.3122
3.36 | -80.3197
3.3604 | -80.3279
3.3608 | -80.3358
3.3611 | -80.3392
3.3615 | -80.3473
3.3619 | -80.3539
3.3622 | -80.3613
3.3626 | -80.3687
3.3629 | -80.3759
3.3633 | -80.3837
3.3637 | -80.3924
3.364 | -80.3998
3.3644 | -80.4063
3.3647 | -80.4148
3.3651 | -80.4222
3.3655 | -80.4308
3.3658 | -80.4403
3.3662 | -80.4469
3.3666 | -80.4551
3.3669 | -80.4651
3.3673 | -80.4747
3.3676 | -80.483
3.368 | -80.4891
3.3684 | -80.4957
3.3687 | -80.5067
3.3691 | -80.5176
3.3695 | -80.5263
3.3698 | -80.5344
3.3702 | -80.5394
3.3705 | -80.5479
3.3709 | -80.56
3.3713 | -80.57
3.3716 | -80.573
3.372 | -80.5891
3.3723 | -80.6022
3.3727 | -80.607
3.3731 | -80.6142
3.3734 | -80.621
3.3738 | -80.6269
3.3742 | -80.6359
3.3745 | -80.6465
3.3749 | -80.6519
3.3752 | -80.6555
3.3756 | -80.6613
3.376 | -80.67
3.3763 | -80.6743
3.3767 | -80.678
3.3771 | -80.6821
3.3774 | -80.6861
3.3778 | -80.6927
3.3781 | -80.6974
3.3785 | -80.7006
3.3789 | -80.704
3.3792 | -80.7073
3.3796 | -80.7133
3.3799 | -80.7224
3.3803 | -80.7275
3.3807 | -80.732
3.381 | -80.7418
3.3814 | -80.7551
3.3818 | -80.7592
3.3821 | -80.7614
3.3825 | -80.7665
3.3828 | -80.7721
3.3832 | -80.7784
3.3836 | -80.7848
3.3839 | -80.794
3.3843 | -80.7979
3.3846 | -80.8029
3.385 | -80.8081
3.3854 | -80.8159
3.3857 | -80.8218
3.3861 | -80.8321
3.3865 | -80.8365
3.3868 | -80.839
3.3872 | -80.8443
3.3875 | -80.8491
3.3879 | -80.8527
3.3883 | -80.8611
3.3886 | -80.871
3.389 | -80.8762
3.3894 | -80.8803
3.3897 | -80.8832
3.3901 | -80.8849
3.3904 | -80.8925
3.3908 | -80.9059
3.3912 | -80.9142
3.3915 | -80.9151
3.3919 | -80.9177
3.3922 | -80.9187
3.3926 | -80.9193
3.393 | -80.9226
3.3933 | -80.9324
3.3937 | -80.9356
3.3941 | -80.9381
3.3944 | -80.9392
3.3948 | -80.9404
3.3951 | -80.9419
3.3955 | -80.9423
3.3959 | -80.9439
3.3962 | -80.9451
3.3966 | -80.9455
3.3969 | -80.9464
3.3973 | -80.9506
3.3977 | -80.9537
3.398 | -80.9549
3.3984 | -80.9555
3.3988 | -80.9566
3.3991 | -80.9563
3.3995 | -80.9563
3.3998 | -80.9584
3.4002 | -80.959
3.4006 | -80.9594
3.4009 | -80.9604
3.4013 | -80.9613
3.4017 | -80.9622
3.402 | -80.9619
3.4024 | -80.9625
3.4027 | -80.9655
3.4031 | -80.9682
3.4035 | -80.9688
3.4038 | -80.9697
3.4042 | -80.9706
3.4045 | -80.9717
3.4049 | -80.9734
3.4053 | -80.9761
3.4056 | -80.9766
3.406 | -80.9762
3.4064 | -80.9777
3.4067 | -80.9797
3.4071 | -80.981
3.4074 | -80.9827
3.4078 | -80.9824
3.4082 | -80.9832
3.4085 | -80.9843
3.4089 | -80.9858
3.4093 | -80.9874
3.4096 | -80.9891
3.41 | -80.9906
3.4103 | -80.9916
3.4107 | -80.9932
3.4111 | -80.9946
3.4114 | -80.9967
3.4118 | -80.9995
3.4121 | -81.0013
3.4125 | -81.0033
3.4129 | -81.0046
3.4132 | -81.0079
3.4136 | -81.0089
3.414 | -81.0107
3.4143 | -81.0119
3.4147 | -81.0124
3.415 | -81.0131
3.4154 | -81.0144
3.4158 | -81.0186
3.4161 | -81.0216
3.4165 | -81.0227
3.4168 | -81.0227
3.4172 | -81.0236
3.4176 | -81.0264
3.4179 | -81.0289
3.4183 | -81.0326
3.4187 | -81.0344
3.419 | -81.0358
3.4194 | -81.0367
3.4197 | -81.0387
3.4201 | -81.0403
3.4205 | -81.0423
3.4208 | -81.043
3.4212 | -81.0422
3.4216 | -81.0419
3.4219 | -81.0438
3.4223 | -81.0468
3.4226 | -81.0484
3.423 | -81.0485
3.4234 | -81.0484
3.4237 | -81.0507
3.4241 | -81.0524
3.4244 | -81.0538
3.4248 | -81.053
3.4252 | -81.05
3.4255 | -81.0509
3.4259 | -81.0526
3.4263 | -81.0546
3.4266 | -81.0563
3.427 | -81.0584
3.4273 | -81.0605
3.4277 | -81.0626
3.4281 | -81.064
3.4284 | -81.0644
3.4288 | -81.0654
3.4291 | -81.0667
3.4295 | -81.0681
3.4299 | -81.0695
3.4302 | -81.0716
3.4306 | -81.0737
3.431 | -81.0753
3.4313 | -81.0794
3.4317 | -81.0836
3.432 | -81.0861
3.4324 | -81.0873
3.4328 | -81.0877
3.4331 | -81.0904
3.4335 | -81.0933
3.4339 | -81.0953
3.4342 | -81.0984
3.4346 | -81.1005
3.4349 | -81.103
3.4353 | -81.1066
3.4357 | -81.1068
3.436 | -81.1083
3.4364 | -81.1105
3.4367 | -81.1127
3.4371 | -81.1144
3.4375 | -81.1167
3.4378 | -81.1186
3.4382 | -81.1202
3.4386 | -81.1215
3.4389 | -81.1235
3.4393 | -81.1259
3.4396 | -81.1269
3.44 | -81.129
3.4404 | -81.131
3.4407 | -81.1324
3.4411 | -81.1347
3.4415 | -81.1395
3.4418 | -81.1438
3.4422 | -81.1472
3.4425 | -81.1511
3.4429 | -81.1547
3.4433 | -81.1561
3.4436 | -81.1574
3.444 | -81.1592
3.4443 | -81.1614
3.4447 | -81.1606
3.4451 | -81.1634
3.4454 | -81.1657
3.4458 | -81.1685
3.4462 | -81.1708
3.4465 | -81.1728
3.4469 | -81.1704
3.4472 | -81.1627
3.4476 | -81.1613
3.448 | -81.1639
3.4483 | -81.1654
3.4487 | -81.1674
3.449 | -81.168
3.4494 | -81.1687
3.4498 | -81.1698
3.4501 | -81.1717
3.4505 | -81.1723
3.4509 | -81.1719
3.4512 | -81.1715
3.4516 | -81.1727
3.4519 | -81.1761
3.4523 | -81.1776
3.4527 | -81.1775
3.453 | -81.1784
3.4534 | -81.1795
3.4538 | -81.1812
3.4541 | -81.182
3.4545 | -81.1831
3.4548 | -81.1836
3.4552 | -81.1852
3.4556 | -81.187
3.4559 | -81.1896
3.4563 | -81.1914
3.4566 | -81.1929
3.457 | -81.1918
3.4574 | -81.1911
3.4577 | -81.1923
3.4581 | -81.1938
3.4585 | -81.1965
3.4588 | -81.1994
3.4592 | -81.2024
3.4595 | -81.2034
3.4599 | -81.2058
3.4603 | -81.208
3.4606 | -81.2098
3.461 | -81.2116
3.4614 | -81.211
3.4617 | -81.209
3.4621 | -81.2065
3.4624 | -81.208
3.4628 | -81.2073
3.4632 | -81.2056
3.4635 | -81.205
3.4639 | -81.206
3.4642 | -81.2017
3.4646 | -81.199
3.465 | -81.2007
3.4653 | -81.2027
3.4657 | -81.2011
3.4661 | -81.2014
3.4664 | -81.2013
3.4668 | -81.2006
3.4671 | -81.1999
3.4675 | -81.1972
3.4679 | -81.1971
3.4682 | -81.1995
3.4686 | -81.2008
3.4689 | -81.2028
3.4693 | -81.2043
3.4697 | -81.2046
3.47 | -81.2054
3.4704 | -81.2074
3.4708 | -81.2082
3.4711 | -81.208
3.4715 | -81.2107
3.4718 | -81.2118
3.4722 | -81.2117
3.4726 | -81.211
3.4729 | -81.21
3.4733 | -81.2103
3.4737 | -81.2114
3.474 | -81.2116
3.4744 | -81.2107
3.4747 | -81.2079
3.4751 | -81.2037
3.4755 | -81.2027
3.4758 | -81.202
3.4762 | -81.2006
3.4765 | -81.1999
3.4769 | -81.2
3.4773 | -81.2005
3.4776 | -81.2006
3.478 | -81.2007
3.4784 | -81.2009
3.4787 | -81.1999
3.4791 | -81.2003
3.4794 | -81.2024
3.4798 | -81.2006
3.4802 | -81.1977
3.4805 | -81.1955
3.4809 | -81.1946
3.4813 | -81.1929
3.4816 | -81.1907
3.482 | -81.1898
3.4823 | -81.189
3.4827 | -81.1879
3.4831 | -81.1873
3.4834 | -81.1858
3.4838 | -81.1841
3.4841 | -81.1831
3.4845 | -81.1816
3.4849 | -81.1781
3.4852 | -81.1766
3.4856 | -81.1752
3.486 | -81.1736
3.4863 | -81.1714
3.4867 | -81.1701
3.487 | -81.1682
3.4874 | -81.1686
3.4878 | -81.1688
3.4881 | -81.1675
3.4885 | -81.1649
3.4888 | -81.1574
3.4892 | -81.1532
3.4896 | -81.1535
3.4899 | -81.1524
3.4903 | -81.1484
3.4907 | -81.1457
3.491 | -81.145
3.4914 | -81.1414
3.4917 | -81.1399
3.4921 | -81.1382
3.4925 | -81.1364
3.4928 | -81.1336
3.4932 | -81.1305
3.4936 | -81.129
3.4939 | -81.1277
3.4943 | -81.1255
3.4946 | -81.1233
3.495 | -81.1197
3.4954 | -81.1171
3.4957 | -81.1148
3.4961 | -81.1132
3.4964 | -81.1112
3.4968 | -81.1075
3.4972 | -81.1044
3.4975 | -81.1033
3.4979 | -81.1027
3.4983 | -81.1007
3.4986 | -81.0985
3.499 | -81.0951
3.4993 | -81.0926
3.4997 | -81.0904
3.5001 | -81.0887
3.5004 | -81.0848
3.5008 | -81.0803
3.5011 | -81.0771
3.5015 | -81.0749
3.5019 | -81.0725
3.5022 | -81.0687
3.5026 | -81.0657
3.503 | -81.0629
3.5033 | -81.0601
3.5037 | -81.0575
3.504 | -81.0548
3.5044 | -81.052
3.5048 | -81.0498
3.5051 | -81.0478
3.5055 | -81.0439
3.5059 | -81.0407
3.5062 | -81.0377
3.5066 | -81.0365
3.5069 | -81.0355
3.5073 | -81.0305
3.5077 | -81.0268
3.508 | -81.0252
3.5084 | -81.023
3.5087 | -81.0199
3.5091 | -81.0167
3.5095 | -81.0134
3.5098 | -81.0075
3.5102 | -81.0016
3.5106 | -80.9982
3.5109 | -80.9947
3.5113 | -80.9914
3.5116 | -80.9881
3.512 | -80.9847
3.5124 | -80.9817
3.5127 | -80.9786
3.5131 | -80.9754
3.5135 | -80.9719
3.5138 | -80.9687
3.5142 | -80.9657
3.5145 | -80.9649
3.5149 | -80.9614
3.5153 | -80.9575
3.5156 | -80.9538
3.516 | -80.9498
3.5163 | -80.9461
3.5167 | -80.9444
3.5171 | -80.9405
3.5174 | -80.9386
3.5178 | -80.9342
3.5182 | -80.9285
3.5185 | -80.9237
3.5189 | -80.9204
3.5192 | -80.9167
3.5196 | -80.9122
3.52 | -80.9086
3.5203 | -80.9055
3.5207 | -80.9004
3.521 | -80.8957
3.5214 | -80.8919
3.5218 | -80.8879
3.5221 | -80.8848
3.5225 | -80.8825
3.5229 | -80.8777
3.5232 | -80.874
3.5236 | -80.8725
3.5239 | -80.8697
3.5243 | -80.8658
3.5247 | -80.8617
3.525 | -80.8567
3.5254 | -80.8493
3.5258 | -80.845
3.5261 | -80.8407
3.5265 | -80.8375
3.5268 | -80.8337
3.5272 | -80.8297
3.5276 | -80.8256
3.5279 | -80.822
3.5283 | -80.8188
3.5286 | -80.8145
3.529 | -80.8111
3.5294 | -80.8092
3.5297 | -80.8061
3.5301 | -80.8023
3.5305 | -80.7998
3.5308 | -80.7976
3.5312 | -80.794
3.5315 | -80.7899
3.5319 | -80.7854
3.5323 | -80.7802
3.5326 | -80.7767
3.533 | -80.7754
3.5333 | -80.7733
3.5337 | -80.7707
3.5341 | -80.7677
3.5344 | -80.7641
3.5348 | -80.7592
3.5352 | -80.754
3.5355 | -80.7495
3.5359 | -80.7463
3.5362 | -80.7429
3.5366 | -80.7398
3.537 | -80.736
3.5373 | -80.7312
3.5377 | -80.7268
3.5381 | -80.7214
3.5384 | -80.717
3.5388 | -80.7125
3.5391 | -80.7077
3.5395 | -80.7017
3.5399 | -80.6938
3.5402 | -80.6883
3.5406 | -80.6841
3.5409 | -80.6806
3.5413 | -80.6783
3.5417 | -80.6752
3.542 | -80.6714
3.5424 | -80.6673
3.5428 | -80.6631
3.5431 | -80.6591
3.5435 | -80.6567
3.5438 | -80.6541
3.5442 | -80.6504
3.5446 | -80.6465
3.5449 | -80.6436
3.5453 | -80.6413
3.5457 | -80.6376
3.546 | -80.6337
3.5464 | -80.6299
3.5467 | -80.6264
3.5471 | -80.6228
3.5475 | -80.6178
3.5478 | -80.613
3.5482 | -80.6094
3.5485 | -80.6058
3.5489 | -80.6039
3.5493 | -80.6003
3.5496 | -80.5965
3.55 | -80.5928
3.5504 | -80.5891
3.5507 | -80.5852
3.5511 | -80.5815
3.5514 | -80.578
3.5518 | -80.576
3.5522 | -80.574
3.5525 | -80.5708
3.5529 | -80.5674
3.5532 | -80.5644
3.5536 | -80.5625
3.554 | -80.5591
3.5543 | -80.5566
3.5547 | -80.5536
3.5551 | -80.5506
3.5554 | -80.5477
3.5558 | -80.5454
3.5561 | -80.5406
3.5565 | -80.5358
3.5569 | -80.5326
3.5572 | -80.5296
3.5576 | -80.5271
3.558 | -80.5247
3.5583 | -80.5218
3.5587 | -80.5192
3.559 | -80.5168
3.5594 | -80.5166
3.5598 | -80.5143
3.5601 | -80.5119
3.5605 | -80.5093
3.5608 | -80.5065
3.5612 | -80.5048
3.5616 | -80.5033
3.5619 | -80.5016
3.5623 | -80.5004
3.5627 | -80.4982
3.563 | -80.496
3.5634 | -80.4937
3.5637 | -80.4899
3.5641 | -80.4849
3.5645 | -80.4821
3.5648 | -80.4799
3.5652 | -80.478
3.5655 | -80.4761
3.5659 | -80.475
3.5663 | -80.4719
3.5666 | -80.4666
3.567 | -80.4638
3.5674 | -80.4628
3.5677 | -80.461
3.5681 | -80.4589
3.5684 | -80.4554
3.5688 | -80.4503
3.5692 | -80.4453
3.5695 | -80.4435
3.5699 | -80.4441
3.5703 | -80.4432
3.5706 | -80.4411
3.571 | -80.4389
3.5713 | -80.4373
3.5717 | -80.4359
3.5721 | -80.4342
3.5724 | -80.4324
3.5728 | -80.4298
3.5731 | -80.4276
3.5735 | -80.426
3.5739 | -80.4249
3.5742 | -80.4233
3.5746 | -80.4217
3.575 | -80.4205
3.5753 | -80.4188
3.5757 | -80.4185
3.576 | -80.418
3.5764 | -80.4169
3.5768 | -80.4154
3.5771 | -80.413
3.5775 | -80.4111
3.5779 | -80.4099
3.5782 | -80.4088
3.5786 | -80.4076
3.5789 | -80.407
3.5793 | -80.408
3.5797 | -80.409
3.58 | -80.4081
3.5804 | -80.4099
3.5807 | -80.4088
3.5811 | -80.408
3.5815 | -80.4056
3.5818 | -80.404
3.5822 | -80.4034
3.5826 | -80.4039
3.5829 | -80.4034
3.5833 | -80.4022
3.5836 | -80.4005
3.584 | -80.4006
3.5844 | -80.4006
3.5847 | -80.3996
3.5851 | -80.3992
3.5854 | -80.3985
3.5858 | -80.3982
3.5862 | -80.3975
3.5865 | -80.397
3.5869 | -80.3964
3.5873 | -80.3968
3.5876 | -80.3971
3.588 | -80.3976
3.5883 | -80.3979
3.5887 | -80.3994
3.5891 | -80.3975
3.5894 | -80.3964
3.5898 | -80.3966
3.5902 | -80.397
3.5905 | -80.3968
3.5909 | -80.3993
3.5912 | -80.3993
3.5916 | -80.3991
3.592 | -80.3974
3.5923 | -80.3969
3.5927 | -80.3985
3.593 | -80.4
3.5934 | -80.4004
3.5938 | -80.4011
3.5941 | -80.4021
3.5945 | -80.4036
3.5949 | -80.4048
3.5952 | -80.4041
3.5956 | -80.403
3.5959 | -80.4033
3.5963 | -80.4041
3.5967 | -80.4041
3.597 | -80.4045
3.5974 | -80.4052
3.5978 | -80.4056
3.5981 | -80.4059
3.5985 | -80.4063
3.5988 | -80.4071
3.5992 | -80.4081
3.5996 | -80.4096
3.5999 | -80.4087
3.6003 | -80.4085
3.6006 | -80.413
3.601 | -80.4163
3.6014 | -80.4196
3.6017 | -80.4214
3.6021 | -80.4226
3.6025 | -80.4233
3.6028 | -80.4243
3.6032 | -80.4258
3.6035 | -80.4267
3.6039 | -80.4277
3.6043 | -80.4293
3.6046 | -80.4303
3.605 | -80.4311
3.6053 | -80.4325
3.6057 | -80.434
3.6061 | -80.4355
3.6064 | -80.4375
3.6068 | -80.4389
3.6072 | -80.4391
3.6075 | -80.441
3.6079 | -80.4424
3.6082 | -80.4434
3.6086 | -80.4449
3.609 | -80.4473
3.6093 | -80.4494
3.6097 | -80.4502
3.6101 | -80.4512
3.6104 | -80.4532
3.6108 | -80.4552
3.6111 | -80.4569
3.6115 | -80.4612
3.6119 | -80.4633
3.6122 | -80.4663
3.6126 | -80.4683
3.6129 | -80.4709
3.6133 | -80.4737
3.6137 | -80.475
3.614 | -80.4762
3.6144 | -80.478
3.6148 | -80.4804
3.6151 | -80.4814
3.6155 | -80.483
3.6158 | -80.4857
3.6162 | -80.4883
3.6166 | -80.4911
3.6169 | -80.4928
3.6173 | -80.4951
3.6176 | -80.4977
3.618 | -80.5
3.6184 | -80.5027
3.6187 | -80.5061
3.6191 | -80.5094
3.6195 | -80.5114
3.6198 | -80.514
3.6202 | -80.5189
3.6205 | -80.5231
3.6209 | -80.5257
3.6213 | -80.5283
3.6216 | -80.5306
3.622 | -80.534
3.6224 | -80.537
3.6227 | -80.5396
3.6231 | -80.5422
3.6234 | -80.5451
3.6238 | -80.5473
3.6242 | -80.5483
3.6245 | -80.5503
3.6249 | -80.554
3.6252 | -80.5573
3.6256 | -80.5599
3.626 | -80.5629
3.6263 | -80.5659
3.6267 | -80.5688
3.6271 | -80.5716
3.6274 | -80.5754
3.6278 | -80.5783
3.6281 | -80.5811
3.6285 | -80.5835
3.6289 | -80.5867
3.6292 | -80.5901
3.6296 | -80.5931
3.63 | -80.5952
3.6303 | -80.5985
3.6307 | -80.6015
3.631 | -80.6039
3.6314 | -80.606
3.6318 | -80.6082
3.6321 | -80.6113
3.6325 | -80.6153
3.6328 | -80.6185
3.6332 | -80.6215
3.6336 | -80.6244
3.6339 | -80.6276
3.6343 | -80.6307
3.6347 | -80.6341
3.635 | -80.6376
3.6354 | -80.6407
3.6357 | -80.6436
3.6361 | -80.6477
3.6365 | -80.6518
3.6368 | -80.6564
3.6372 | -80.6608
3.6375 | -80.6648
3.6379 | -80.6692
3.6383 | -80.6724
3.6386 | -80.6755
3.639 | -80.6788
3.6394 | -80.6824
3.6397 | -80.686
3.6401 | -80.6905
3.6404 | -80.6936
3.6408 | -80.6966
3.6412 | -80.7005
3.6415 | -80.7049
3.6419 | -80.7093
3.6423 | -80.7134
3.6426 | -80.717
3.643 | -80.7219
3.6433 | -80.7252
3.6437 | -80.7309
3.6441 | -80.7361
3.6444 | -80.7409
3.6448 | -80.7455
3.6451 | -80.7506
3.6455 | -80.7554
3.6459 | -80.7594
3.6462 | -80.7632
3.6466 | -80.7672
3.647 | -80.7711
3.6473 | -80.775
3.6477 | -80.7786
3.648 | -80.7816
3.6484 | -80.7853
3.6488 | -80.7891
3.6491 | -80.7929
3.6495 | -80.7967
3.6498 | -80.8005
3.6502 | -80.8039
3.6506 | -80.8078
3.6509 | -80.8122
3.6513 | -80.816
3.6517 | -80.8186
3.652 | -80.8228
3.6524 | -80.826
3.6527 | -80.8285
3.6531 | -80.8336
3.6535 | -80.8384
3.6538 | -80.8418
3.6542 | -80.8451
3.6546 | -80.8494
3.6549 | -80.8532
3.6553 | -80.8566
3.6556 | -80.8606
3.656 | -80.8646
3.6564 | -80.8683
3.6567 | -80.8725
3.6571 | -80.8749
3.6574 | -80.8768
3.6578 | -80.8812
3.6582 | -80.8858
3.6585 | -80.8895
3.6589 | -80.8931
3.6593 | -80.8973
3.6596 | -80.9017
3.66 | -80.905
3.6603 | -80.9087
3.6607 | -80.9134
3.6611 | -80.9184
3.6614 | -80.923
3.6618 | -80.9267
3.6622 | -80.9306
3.6625 | -80.9351
3.6629 | -80.9404
3.6632 | -80.9449
3.6636 | -80.9484
3.664 | -80.9531
3.6643 | -80.9563
3.6647 | -80.9601
3.665 | -80.9637
3.6654 | -80.9681
3.6658 | -80.9726
3.6661 | -80.9763
3.6665 | -80.9795
3.6669 | -80.9836
3.6672 | -80.9888
3.6676 | -80.9952
3.6679 | -81.0014
3.6683 | -81.0062
3.6687 | -81.0097
3.669 | -81.0123
3.6694 | -81.0161
3.6697 | -81.0208
3.6701 | -81.0247
3.6705 | -81.0279
3.6708 | -81.0318
3.6712 | -81.0361
3.6716 | -81.042
3.6719 | -81.0501
3.6723 | -81.0542
3.6726 | -81.0577
3.673 | -81.0614
3.6734 | -81.0659
3.6737 | -81.0699
3.6741 | -81.0736
3.6745 | -81.0778
3.6748 | -81.0817
3.6752 | -81.0869
3.6755 | -81.092
3.6759 | -81.0964
3.6763 | -81.1007
3.6766 | -81.1045
3.677 | -81.108
3.6773 | -81.1116
3.6777 | -81.1153
3.6781 | -81.1192
3.6784 | -81.123
3.6788 | -81.1263
3.6792 | -81.1303
3.6795 | -81.1351
3.6799 | -81.1399
3.6802 | -81.1428
3.6806 | -81.1452
3.681 | -81.1486
3.6813 | -81.1519
3.6817 | -81.1559
3.682 | -81.1596
3.6824 | -81.1641
3.6828 | -81.1676
3.6831 | -81.1726
3.6835 | -81.1768
3.6839 | -81.18
3.6842 | -81.1842
3.6846 | -81.1888
3.6849 | -81.1931
3.6853 | -81.1962
3.6857 | -81.2002
3.686 | -81.2043
3.6864 | -81.2079
3.6868 | -81.211
3.6871 | -81.2143
3.6875 | -81.2189
3.6878 | -81.223
3.6882 | -81.2269
3.6886 | -81.2315
3.6889 | -81.2359
3.6893 | -81.2401
3.6896 | -81.2441
3.69 | -81.2482
3.6904 | -81.2517
3.6907 | -81.2556
3.6911 | -81.2601
3.6915 | -81.2634
3.6918 | -81.2674
3.6922 | -81.2696
3.6925 | -81.2724
3.6929 | -81.2771
3.6933 | -81.279
3.6936 | -81.2811
3.694 | -81.2847
3.6944 | -81.2886
3.6947 | -81.292
3.6951 | -81.2954
3.6954 | -81.3017
3.6958 | -81.3133
3.6962 | -81.3191
3.6965 | -81.3223
3.6969 | -81.3262
3.6972 | -81.3295
3.6976 | -81.3318
3.698 | -81.3354
3.6983 | -81.3395
3.6987 | -81.3434
3.6991 | -81.3475
3.6994 | -81.3537
3.6998 | -81.3585
3.7001 | -81.3624
3.7005 | -81.3657
3.7009 | -81.3698
3.7012 | -81.3751
3.7016 | -81.3787
3.7019 | -81.3798
3.7023 | -81.3826
3.7027 | -81.3868
3.703 | -81.3903
3.7034 | -81.3939
3.7038 | -81.3975
3.7041 | -81.4007
3.7045 | -81.4023
3.7048 | -81.4044
3.7052 | -81.4071
3.7056 | -81.4108
3.7059 | -81.4135
3.7063 | -81.4193
3.7067 | -81.4242
3.707 | -81.4274
3.7074 | -81.4305
3.7077 | -81.4339
3.7081 | -81.4377
3.7085 | -81.4419
3.7088 | -81.4473
3.7092 | -81.4503
3.7095 | -81.4526
3.7099 | -81.4558
3.7103 | -81.4601
3.7106 | -81.4656
3.711 | -81.4718
3.7114 | -81.4785
3.7117 | -81.482
3.7121 | -81.4859
3.7124 | -81.4895
3.7128 | -81.4938
3.7132 | -81.4969
3.7135 | -81.5043
3.7139 | -81.5077
3.7142 | -81.5095
3.7146 | -81.5127
3.715 | -81.5165
3.7153 | -81.5197
3.7157 | -81.523
3.7161 | -81.5257
3.7164 | -81.5294
3.7168 | -81.533
3.7171 | -81.5356
3.7175 | -81.5405
3.7179 | -81.5446
3.7182 | -81.5467
3.7186 | -81.5485
3.719 | -81.5492
3.7193 | -81.5506
3.7197 | -81.5534
3.72 | -81.5557
3.7204 | -81.558
3.7208 | -81.5621
3.7211 | -81.5655
3.7215 | -81.5679
3.7218 | -81.5702
3.7222 | -81.5732
3.7226 | -81.5766
3.7229 | -81.5802
3.7233 | -81.5832
3.7237 | -81.5857
3.724 | -81.5886
3.7244 | -81.5926
3.7247 | -81.5984
3.7251 | -81.6048
3.7255 | -81.6091
3.7258 | -81.6114
3.7262 | -81.6125
3.7266 | -81.6156
3.7269 | -81.6195
3.7273 | -81.6222
3.7276 | -81.6249
3.728 | -81.6271
3.7284 | -81.6292
3.7287 | -81.632
3.7291 | -81.6346
3.7294 | -81.6385
3.7298 | -81.6411
3.7302 | -81.6437
3.7305 | -81.6481
3.7309 | -81.6529
3.7313 | -81.6576
3.7316 | -81.6607
3.732 | -81.6646
3.7323 | -81.6673
3.7327 | -81.6685
3.7331 | -81.6711
3.7334 | -81.6741
3.7338 | -81.6769
3.7341 | -81.6803
3.7345 | -81.6837
3.7349 | -81.6855
3.7352 | -81.6889
3.7356 | -81.6942
3.736 | -81.6972
3.7363 | -81.7006
3.7367 | -81.704
3.737 | -81.7066
3.7374 | -81.7086
3.7378 | -81.7111
3.7381 | -81.7142
3.7385 | -81.7168
3.7389 | -81.7194
3.7392 | -81.723
3.7396 | -81.7274
3.7399 | -81.7319
3.7403 | -81.7409
3.7407 | -81.7462
3.741 | -81.7478
3.7414 | -81.7495
3.7417 | -81.7525
3.7421 | -81.7561
3.7425 | -81.7586
3.7428 | -81.764
3.7432 | -81.7669
3.7436 | -81.7697
3.7439 | -81.7733
3.7443 | -81.7759
3.7446 | -81.7771
3.745 | -81.7791
3.7454 | -81.7821
3.7457 | -81.7846
3.7461 | -81.787
3.7464 | -81.7908
3.7468 | -81.793
3.7472 | -81.7952
3.7475 | -81.7893
3.7479 | -81.7881
3.7483 | -81.79
3.7486 | -81.7923
3.749 | -81.7943
3.7493 | -81.7966
3.7497 | -81.799
3.7501 | -81.801
3.7504 | -81.8028
3.7508 | -81.805
3.7512 | -81.8078
3.7515 | -81.8094
3.7519 | -81.8116
3.7522 | -81.8139
3.7526 | -81.8174
3.753 | -81.8202
3.7533 | -81.8207
3.7537 | -81.8217
3.754 | -81.8245
3.7544 | -81.8276
3.7548 | -81.8292
3.7551 | -81.8301
3.7555 | -81.8316
3.7559 | -81.8334
3.7562 | -81.8348
3.7566 | -81.8348
3.7569 | -81.8359
3.7573 | -81.8378
3.7577 | -81.8395
3.758 | -81.8419
3.7584 | -81.8438
3.7588 | -81.8457
3.7591 | -81.8475
3.7595 | -81.8498
3.7598 | -81.8516
3.7602 | -81.8552
3.7606 | -81.8578
3.7609 | -81.8592
3.7613 | -81.8613
3.7616 | -81.8636
3.762 | -81.8651
3.7624 | -81.8674
3.7627 | -81.8731
3.7631 | -81.8774
3.7635 | -81.8788
3.7638 | -81.8802
3.7642 | -81.8818
3.7645 | -81.8831
3.7649 | -81.8835
3.7653 | -81.8847
3.7656 | -81.8871
3.766 | -81.8888
3.7663 | -81.8902
3.7667 | -81.8915
3.7671 | -81.8929
3.7674 | -81.8939
3.7678 | -81.8952
3.7682 | -81.8956
3.7685 | -81.8945
3.7689 | -81.8948
3.7692 | -81.897
3.7696 | -81.8996
3.77 | -81.9009
3.7703 | -81.9013
3.7707 | -81.901
3.7711 | -81.901
3.7714 | -81.9032
3.7718 | -81.9077
3.7721 | -81.9076
3.7725 | -81.9072
3.7729 | -81.9085
3.7732 | -81.9104
3.7736 | -81.9123
3.7739 | -81.9145
3.7743 | -81.9186
3.7747 | -81.9221
3.775 | -81.9241
3.7754 | -81.9254
3.7758 | -81.9264
3.7761 | -81.9269
3.7765 | -81.9293
3.7768 | -81.9306
3.7772 | -81.9304
3.7776 | -81.9311
3.7779 | -81.9329
3.7783 | -81.9341
3.7786 | -81.9349
3.779 | -81.9363
3.7794 | -81.9391
3.7797 | -81.9408
3.7801 | -81.9418
3.7805 | -81.9431
3.7808 | -81.9448
3.7812 | -81.9451
3.7815 | -81.9448
3.7819 | -81.9444
3.7823 | -81.9447
3.7826 | -81.9456
3.783 | -81.948
3.7834 | -81.9489
3.7837 | -81.9494
3.7841 | -81.9502
3.7844 | -81.9509
3.7848 | -81.9515
3.7852 | -81.9526
3.7855 | -81.9536
3.7859 | -81.9544
3.7862 | -81.9558
3.7866 | -81.956
3.787 | -81.9534
3.7873 | -81.9508
3.7877 | -81.9509
3.7881 | -81.9524
3.7884 | -81.952
3.7888 | -81.9519
3.7891 | -81.9525
3.7895 | -81.9542
3.7899 | -81.9559
3.7902 | -81.9564
3.7906 | -81.9567
3.791 | -81.9573
3.7913 | -81.9572
3.7917 | -81.9576
3.792 | -81.9576
3.7924 | -81.9563
3.7928 | -81.9566
3.7931 | -81.958
3.7935 | -81.9588
3.7938 | -81.9596
3.7942 | -81.9609
3.7946 | -81.9634
3.7949 | -81.9687
3.7953 | -81.9713
3.7957 | -81.9718
3.796 | -81.9709
3.7964 | -81.9711
3.7967 | -81.9711
3.7971 | -81.9713
3.7975 | -81.9716
3.7978 | -81.972
3.7982 | -81.9726
3.7985 | -81.9721
3.7989 | -81.9713
3.7993 | -81.9703
3.7996 | -81.9698
3.8 | -81.9697
3.8004 | -81.9695
3.8007 | -81.9686
3.8011 | -81.9682
3.8014 | -81.9682
3.8018 | -81.9684
3.8022 | -81.9691
3.8025 | -81.969
3.8029 | -81.9705
3.8033 | -81.9718
3.8036 | -81.9718
3.804 | -81.972
3.8043 | -81.9723
3.8047 | -81.972
3.8051 | -81.9715
3.8054 | -81.971
3.8058 | -81.9715
3.8061 | -81.9728
3.8065 | -81.9728
3.8069 | -81.9733
3.8072 | -81.9734
3.8076 | -81.9733
3.808 | -81.9732
3.8083 | -81.9726
3.8087 | -81.9721
3.809 | -81.9724
3.8094 | -81.9718
3.8098 | -81.9716
3.8101 | -81.9725
3.8105 | -81.9724
3.8108 | -81.9726
3.8112 | -81.9725
3.8116 | -81.9724
3.8119 | -81.9726
3.8123 | -81.975
3.8127 | -81.9782
3.813 | -81.979
3.8134 | -81.979
3.8137 | -81.9789
3.8141 | -81.9793
3.8145 | -81.9789
3.8148 | -81.9782
3.8152 | -81.9785
3.8156 | -81.9785
3.8159 | -81.978
3.8163 | -81.9763
3.8166 | -81.9753
3.817 | -81.9756
3.8174 | -81.9749
3.8177 | -81.974
3.8181 | -81.975
3.8184 | -81.9761
3.8188 | -81.9747
3.8192 | -81.9738
3.8195 | -81.9747
3.8199 | -81.9747
3.8203 | -81.973
3.8206 | -81.9746
3.821 | -81.9771
3.8213 | -81.9746
3.8217 | -81.9736
3.8221 | -81.9738
3.8224 | -81.9743
3.8228 | -81.9729
3.8232 | -81.973
3.8235 | -81.9764
3.8239 | -81.9772
3.8242 | -81.9772
3.8246 | -81.977
3.825 | -81.9777
3.8253 | -81.9793
3.8257 | -81.9833
3.826 | -81.9876
3.8264 | -81.9884
3.8268 | -81.9868
3.8271 | -81.9852
3.8275 | -81.9847
3.8279 | -81.984
3.8282 | -81.9835
3.8286 | -81.9832
3.8289 | -81.9829
3.8293 | -81.9824
3.8297 | -81.9822
3.83 | -81.9828
3.8304 | -81.9825
3.8307 | -81.9828
3.8311 | -81.9829
3.8315 | -81.9819
3.8318 | -81.981
3.8322 | -81.981
3.8326 | -81.9812
3.8329 | -81.9815
3.8333 | -81.9816
3.8336 | -81.981
3.834 | -81.9805
3.8344 | -81.9809
3.8347 | -81.9825
3.8351 | -81.9835
3.8355 | -81.9834
3.8358 | -81.9832
3.8362 | -81.9834
3.8365 | -81.9827
3.8369 | -81.9826
3.8373 | -81.9828
3.8376 | -81.9828
3.838 | -81.9832
3.8383 | -81.9825
3.8387 | -81.9794
3.8391 | -81.98
3.8394 | -81.9827
3.8398 | -81.9833
3.8402 | -81.9828
3.8405 | -81.9831
3.8409 | -81.9833
3.8412 | -81.9848
3.8416 | -81.985
3.842 | -81.9865
3.8423 | -81.986
3.8427 | -81.9854
3.8431 | -81.9873
3.8434 | -81.9881
3.8438 | -81.9889
3.8441 | -81.9892
3.8445 | -81.9897
3.8449 | -81.99
3.8452 | -81.9903
3.8456 | -81.9889
3.8459 | -81.9879
3.8463 | -81.9864
3.8467 | -81.9867
3.847 | -81.987
3.8474 | -81.9907
3.8478 | -81.9921
3.8481 | -81.9927
3.8485 | -81.9931
3.8488 | -81.9934
3.8492 | -81.9945
3.8496 | -81.9941
3.8499 | -81.9927
3.8503 | -81.9891
3.8506 | -81.99
3.851 | -81.991
3.8514 | -81.9895
3.8517 | -81.9898
3.8521 | -81.9882
3.8525 | -81.9864
3.8528 | -81.9868
3.8532 | -81.9867
3.8535 | -81.9868
3.8539 | -81.9874
3.8543 | -81.9894
3.8546 | -81.992
3.855 | -81.993
3.8554 | -81.992
3.8557 | -81.9925
3.8561 | -81.9931
3.8564 | -81.9934
3.8568 | -81.9942
3.8572 | -81.9952
3.8575 | -81.9956
3.8579 | -81.9955
3.8582 | -81.996
3.8586 | -81.9965
3.859 | -81.9989
3.8593 | -82.0017
3.8597 | -82.0027
3.8601 | -82.0031
3.8604 | -82.0038
3.8608 | -82.0042
3.8611 | -82.0043
3.8615 | -82.0045
3.8619 | -82.0054
3.8622 | -82.0074
3.8626 | -82.0079
3.863 | -82.0086
3.8633 | -82.0094
3.8637 | -82.0095
3.864 | -82.0103
3.8644 | -82.0109
3.8648 | -82.0116
3.8651 | -82.0112
3.8655 | -82.011
3.8658 | -82.01
3.8662 | -82.0089
3.8666 | -82.0106
3.8669 | -82.0128
3.8673 | -82.0126
3.8677 | -82.0109
3.868 | -82.0102
3.8684 | -82.0099
3.8687 | -82.0104
3.8691 | -82.0099
3.8695 | -82.0096
3.8698 | -82.0096
3.8702 | -82.0079
3.8705 | -82.0054
3.8709 | -82.0048
3.8713 | -82.0057
3.8716 | -82.0054
3.872 | -82.0055
3.8724 | -82.0053
3.8727 | -82.0049
3.8731 | -82.0043
3.8734 | -82.005
3.8738 | -82.0055
3.8742 | -82.0072
3.8745 | -82.0087
3.8749 | -82.0096
3.8753 | -82.0097
3.8756 | -82.0117
3.876 | -82.0144
3.8763 | -82.0152
3.8767 | -82.0157
3.8771 | -82.0164
3.8774 | -82.0172
3.8778 | -82.0175
3.8781 | -82.0183
3.8785 | -82.019
3.8789 | -82.0194
3.8792 | -82.0192
3.8796 | -82.0196
3.88 | -82.02
3.8803 | -82.0205
3.8807 | -82.0203
3.881 | -82.0209
3.8814 | -82.0203
3.8818 | -82.0208
3.8821 | -82.0216
3.8825 | -82.0206
3.8828 | -82.0206
3.8832 | -82.0208
3.8836 | -82.0217
3.8839 | -82.023
3.8843 | -82.0236
3.8847 | -82.0232
3.885 | -82.0221
3.8854 | -82.024
3.8857 | -82.0257
3.8861 | -82.0251
3.8865 | -82.0232
3.8868 | -82.0242
3.8872 | -82.0247
3.8876 | -82.0246
3.8879 | -82.0245
3.8883 | -82.0252
3.8886 | -82.0254
3.889 | -82.0266
3.8894 | -82.0271
3.8897 | -82.029
3.8901 | -82.0295
3.8904 | -82.03
3.8908 | -82.0314
3.8912 | -82.0333
3.8915 | -82.0338
3.8919 | -82.0343
3.8923 | -82.0365
3.8926 | -82.0376
3.893 | -82.0364
3.8933 | -82.035
3.8937 | -82.0359
3.8941 | -82.0367
3.8944 | -82.0372
3.8948 | -82.0369
3.8952 | -82.0374
3.8955 | -82.0378
3.8959 | -82.0379
3.8962 | -82.0386
3.8966 | -82.0387
3.897 | -82.0392
3.8973 | -82.04
3.8977 | -82.0414
3.898 | -82.0412
3.8984 | -82.0402
3.8988 | -82.0405
3.8991 | -82.0402
3.8995 | -82.0405
3.8999 | -82.0409
3.9002 | -82.0403
3.9006 | -82.0391
3.9009 | -82.0377
3.9013 | -82.0377
3.9017 | -82.0381
3.902 | -82.038
3.9024 | -82.0368
3.9027 | -82.0364
3.9031 | -82.0374
3.9035 | -82.0379
3.9038 | -82.0383
3.9042 | -82.0374
3.9046 | -82.0374
3.9049 | -82.0378
3.9053 | -82.0373
3.9056 | -82.0347
3.906 | -82.0346
3.9064 | -82.0354
3.9067 | -82.0356
3.9071 | -82.0353
3.9075 | -82.0351
3.9078 | -82.0356
3.9082 | -82.0357
3.9085 | -82.0359
3.9089 | -82.0364
3.9093 | -82.0366
3.9096 | -82.038
3.91 | -82.04
3.9103 | -82.0406
3.9107 | -82.041
3.9111 | -82.0406
3.9114 | -82.0403
3.9118 | -82.0395
3.9122 | -82.04
3.9125 | -82.0402
3.9129 | -82.0411
3.9132 | -82.0413
3.9136 | -82.042
3.914 | -82.0437
3.9143 | -82.0442
3.9147 | -82.0462
3.915 | -82.0457
3.9154 | -82.0467
3.9158 | -82.0484
3.9161 | -82.0487
3.9165 | -82.0481
3.9169 | -82.048
3.9172 | -82.0488
3.9176 | -82.0469
3.9179 | -82.0454
3.9183 | -82.0459
3.9187 | -82.0464
3.919 | -82.0465
3.9194 | -82.0451
3.9198 | -82.0436
3.9201 | -82.0424
3.9205 | -82.0438
3.9208 | -82.0465
3.9212 | -82.0464
3.9216 | -82.0466
3.9219 | -82.0474
3.9223 | -82.0476
3.9226 | -82.0476
3.923 | -82.0473
3.9234 | -82.0473
3.9237 | -82.0481
3.9241 | -82.0483
3.9245 | -82.0486
3.9248 | -82.0453
3.9252 | -82.0443
3.9255 | -82.0438
3.9259 | -82.0441
3.9263 | -82.0445
3.9266 | -82.044
3.927 | -82.0437
3.9274 | -82.0443
3.9277 | -82.043
3.9281 | -82.0428
3.9284 | -82.0437
3.9288 | -82.0439
3.9292 | -82.0447
3.9295 | -82.045
3.9299 | -82.0447
3.9302 | -82.0442
3.9306 | -82.0442
3.931 | -82.0441
3.9313 | -82.0437
3.9317 | -82.0438
3.9321 | -82.0435
3.9324 | -82.043
3.9328 | -82.0429
3.9331 | -82.0448
3.9335 | -82.0471
3.9339 | -82.048
3.9342 | -82.0478
3.9346 | -82.049
3.9349 | -82.0496
3.9353 | -82.0501
3.9357 | -82.0504
3.936 | -82.0509
3.9364 | -82.0506
3.9368 | -82.0511
3.9371 | -82.0512
3.9375 | -82.0512
3.9378 | -82.051
3.9382 | -82.0507
3.9386 | -82.0505
3.9389 | -82.0507
3.9393 | -82.0501
3.9397 | -82.0494
3.94 | -82.0492
3.9404 | -82.0491
3.9407 | -82.0484
3.9411 | -82.0479
3.9415 | -82.048
3.9418 | -82.0483
3.9422 | -82.0474
3.9425 | -82.0472
3.9429 | -82.0468
3.9433 | -82.047
3.9436 | -82.0474
3.944 | -82.0472
3.9444 | -82.0457
3.9447 | -82.0452
3.9451 | -82.0461
3.9454 | -82.0471
3.9458 | -82.0456
3.9462 | -82.0461
3.9465 | -82.046
3.9469 | -82.046
3.9472 | -82.0457
3.9476 | -82.045
3.948 | -82.0448
3.9483 | -82.0453
3.9487 | -82.0456
3.9491 | -82.0465
3.9494 | -82.0459
3.9498 | -82.0445
3.9501 | -82.0438
3.9505 | -82.0436
3.9509 | -82.0433
3.9512 | -82.0418
3.9516 | -82.0437
3.952 | -82.0447
3.9523 | -82.045
3.9527 | -82.0457
3.953 | -82.0455
3.9534 | -82.0452
3.9538 | -82.0452
3.9541 | -82.0451
3.9545 | -82.0447
3.9548 | -82.0444
3.9552 | -82.0441
3.9556 | -82.0433
3.9559 | -82.0429
3.9563 | -82.0429
3.9567 | -82.043
3.957 | -82.0433
3.9574 | -82.043
3.9577 | -82.0416
3.9581 | -82.0433
3.9585 | -82.0444
3.9588 | -82.0452
3.9592 | -82.0468
3.9596 | -82.0487
3.9599 | -82.0493
3.9603 | -82.0486
3.9606 | -82.0478
3.961 | -82.0473
3.9614 | -82.0475
3.9617 | -82.0489
3.9621 | -82.0489
3.9624 | -82.0488
3.9628 | -82.0481
3.9632 | -82.0474
3.9635 | -82.0458
3.9639 | -82.0465
3.9643 | -82.0472
3.9646 | -82.0467
3.965 | -82.0456
3.9653 | -82.0436
3.9657 | -82.0412
3.9661 | -82.0399
3.9664 | -82.0376
3.9668 | -82.0363
3.9671 | -82.0356
3.9675 | -82.0352
3.9679 | -82.0342
3.9682 | -82.0332
3.9686 | -82.0354
3.969 | -82.0353
3.9693 | -82.0342
3.9697 | -82.0328
3.97 | -82.0314
3.9704 | -82.0308
3.9708 | -82.0303
3.9711 | -82.0315
3.9715 | -82.0323
3.9719 | -82.0318
3.9722 | -82.0303
3.9726 | -82.0282
3.9729 | -82.0279
3.9733 | -82.0263
3.9737 | -82.0252
3.974 | -82.0248
3.9744 | -82.0252
3.9747 | -82.0247
3.9751 | -82.024
3.9755 | -82.0222
3.9758 | -82.0203
3.9762 | -82.0202
3.9766 | -82.0209
3.9769 | -82.0216
3.9773 | -82.022
3.9776 | -82.0214
3.978 | -82.0214
3.9784 | -82.0215
3.9787 | -82.0215
3.9791 | -82.0209
3.9794 | -82.0213
3.9798 | -82.0231
3.9802 | -82.0235
3.9805 | -82.0228
3.9809 | -82.0235
3.9813 | -82.0266
3.9816 | -82.0276
3.982 | -82.0269
3.9823 | -82.0274
3.9827 | -82.0272
3.9831 | -82.027
3.9834 | -82.0274
3.9838 | -82.0276
3.9842 | -82.0274
3.9845 | -82.0276
3.9849 | -82.028
3.9852 | -82.029
3.9856 | -82.0285
3.986 | -82.0291
3.9863 | -82.0305
3.9867 | -82.0324
3.987 | -82.0339
3.9874 | -82.0342
3.9878 | -82.0327
3.9881 | -82.0311
3.9885 | -82.0308
3.9889 | -82.0309
3.9892 | -82.031
3.9896 | -82.031
3.9899 | -82.0305
3.9903 | -82.0293
3.9907 | -82.0282
3.991 | -82.0277
3.9914 | -82.0272
3.9918 | -82.0256
3.9921 | -82.0243
3.9925 | -82.0237
3.9928 | -82.0236
3.9932 | -82.0242
3.9936 | -82.0245
3.9939 | -82.0242
3.9943 | -82.0255
3.9946 | -82.0256
3.995 | -82.0254
3.9954 | -82.0246
3.9957 | -82.0242
3.9961 | -82.0239
3.9965 | -82.0229
3.9968 | -82.0209
3.9972 | -82.0186
3.9975 | -82.0178
3.9979 | -82.0191
3.9983 | -82.0187
3.9986 | -82.0199
3.999 | -82.0203
3.9993 | -82.0198
3.9997 | -82.0195
4.0001 | -82.0182
4.0004 | -82.0183
4.0008 | -82.0187
4.0012 | -82.0211
4.0015 | -82.0216
4.0019 | -82.0206
4.0022 | -82.0206
4.0026 | -82.0208
4.003 | -82.021
4.0033 | -82.0215
4.0037 | -82.0227
4.0041 | -82.022
4.0044 | -82.0205
4.0048 | -82.0204
4.0051 | -82.0204
4.0055 | -82.0203
4.0059 | -82.0206
4.0062 | -82.0197
4.0066 | -82.02
4.0069 | -82.0193
4.0073 | -82.019
4.0077 | -82.0195
4.008 | -82.02
4.0084 | -82.0218
4.0088 | -82.0218
4.0091 | -82.0221
4.0095 | -82.0224
4.0098 | -82.0225
4.0102 | -82.0226
4.0106 | -82.0237
4.0109 | -82.0251
4.0113 | -82.0263
4.0117 | -82.0264
4.012 | -82.0266
4.0124 | -82.0266
4.0127 | -82.0267
4.0131 | -82.0265
4.0135 | -82.0269
4.0138 | -82.0276
4.0142 | -82.0297
4.0145 | -82.0302
4.0149 | -82.0306
4.0153 | -82.0315
4.0156 | -82.0319
4.016 | -82.0325
4.0164 | -82.0333
4.0167 | -82.0335
4.0171 | -82.0367
4.0174 | -82.0426
4.0178 | -82.0465
4.0182 | -82.0474
4.0185 | -82.0476
4.0189 | -82.0498
4.0192 | -82.0509
4.0196 | -82.0514
4.02 | -82.0521
4.0203 | -82.0525
4.0207 | -82.0527
4.0211 | -82.0536
4.0214 | -82.054
4.0218 | -82.0543
4.0221 | -82.0542
4.0225 | -82.0546
4.0229 | -82.0551
4.0232 | -82.0568
4.0236 | -82.0591
4.024 | -82.0603
4.0243 | -82.0611
4.0247 | -82.0617
4.025 | -82.0627
4.0254 | -82.0644
4.0258 | -82.0653
4.0261 | -82.0663
4.0265 | -82.068
4.0268 | -82.0706
4.0272 | -82.072
4.0276 | -82.073
4.0279 | -82.0739
4.0283 | -82.0746
4.0287 | -82.076
4.029 | -82.0766
4.0294 | -82.0769
4.0297 | -82.078
4.0301 | -82.0787
4.0305 | -82.0785
4.0308 | -82.0783
4.0312 | -82.0788
4.0315 | -82.0797
4.0319 | -82.0814
4.0323 | -82.0817
4.0326 | -82.0817
4.033 | -82.083
4.0334 | -82.0842
4.0337 | -82.0849
4.0341 | -82.0862
4.0344 | -82.0875
4.0348 | -82.0889
4.0352 | -82.0898
4.0355 | -82.0903
4.0359 | -82.0918
4.0363 | -82.093
4.0366 | -82.094
4.037 | -82.0947
4.0373 | -82.0956
4.0377 | -82.0965
4.0381 | -82.0981
4.0384 | -82.0993
4.0388 | -82.1009
4.0391 | -82.1009
4.0395 | -82.0996
4.0399 | -82.0984
4.0402 | -82.1023
4.0406 | -82.1049
4.041 | -82.1056
4.0413 | -82.1068
4.0417 | -82.1077
4.042 | -82.1095
4.0424 | -82.1121
4.0428 | -82.1138
4.0431 | -82.1149
4.0435 | -82.1165
4.0438 | -82.1202
4.0442 | -82.1229
4.0446 | -82.1237
4.0449 | -82.1247
4.0453 | -82.1258
4.0457 | -82.1257
4.046 | -82.1254
4.0464 | -82.1266
4.0467 | -82.1279
4.0471 | -82.1292
4.0475 | -82.1304
4.0478 | -82.1325
4.0482 | -82.134
4.0486 | -82.1361
4.0489 | -82.1383
4.0493 | -82.1394
4.0496 | -82.1404
4.05 | -82.143
4.0504 | -82.1444
4.0507 | -82.1458
4.0511 | -82.1477
4.0514 | -82.1498
4.0518 | -82.1524
4.0522 | -82.1531
4.0525 | -82.154
4.0529 | -82.1545
4.0533 | -82.1553
4.0536 | -82.1574
4.054 | -82.1594
4.0543 | -82.1615
4.0547 | -82.163
4.0551 | -82.1644
4.0554 | -82.1655
4.0558 | -82.1665
4.0562 | -82.1677
4.0565 | -82.1689
4.0569 | -82.1703
4.0572 | -82.172
4.0576 | -82.1728
4.058 | -82.173
4.0583 | -82.1735
4.0587 | -82.1742
4.059 | -82.1777
4.0594 | -82.181
4.0598 | -82.185
4.0601 | -82.1873
4.0605 | -82.1886
4.0609 | -82.19
4.0612 | -82.1923
4.0616 | -82.1941
4.0619 | -82.1949
4.0623 | -82.1963
4.0627 | -82.1976
4.063 | -82.1989
4.0634 | -82.2011
4.0637 | -82.203
4.0641 | -82.2038
4.0645 | -82.2051
4.0648 | -82.2067
4.0652 | -82.2085
4.0656 | -82.2125
4.0659 | -82.2166
4.0663 | -82.2181
4.0666 | -82.2191
4.067 | -82.2195
4.0674 | -82.2206
4.0677 | -82.2213
4.0681 | -82.2228
4.0685 | -82.2239
4.0688 | -82.2256
4.0692 | -82.2285
4.0695 | -82.2302
4.0699 | -82.2313
4.0703 | -82.2321
4.0706 | -82.2332
4.071 | -82.2357
4.0713 | -82.2371
4.0717 | -82.2377
4.0721 | -82.2388
4.0724 | -82.2411
4.0728 | -82.2441
4.0732 | -82.2467
4.0735 | -82.2481
4.0739 | -82.2488
4.0742 | -82.2502
4.0746 | -82.2531
4.075 | -82.2557
4.0753 | -82.257
4.0757 | -82.2583
4.0761 | -82.2589
4.0764 | -82.2607
4.0768 | -82.2625
4.0771 | -82.2627
4.0775 | -82.264
4.0779 | -82.2649
4.0782 | -82.2667
4.0786 | -82.268
4.0789 | -82.2694
4.0793 | -82.27
4.0797 | -82.2708
4.08 | -82.2723
4.0804 | -82.2744
4.0808 | -82.2761
4.0811 | -82.2776
4.0815 | -82.2795
4.0818 | -82.28
4.0822 | -82.2812
4.0826 | -82.283
4.0829 | -82.2848
4.0833 | -82.2849
4.0836 | -82.2871
4.084 | -82.2888
4.0844 | -82.2902
4.0847 | -82.2917
4.0851 | -82.2931
4.0855 | -82.2947
4.0858 | -82.2961
4.0862 | -82.2977
4.0865 | -82.2984
4.0869 | -82.2998
4.0873 | -82.3015
4.0876 | -82.3031
4.088 | -82.3044
4.0884 | -82.3061
4.0887 | -82.3073
4.0891 | -82.3091
4.0894 | -82.3111
4.0898 | -82.3119
4.0902 | -82.3135
4.0905 | -82.3154
4.0909 | -82.3176
4.0912 | -82.3177
4.0916 | -82.3182
4.092 | -82.3195
4.0923 | -82.322
4.0927 | -82.3237
4.0931 | -82.3251
4.0934 | -82.3267
4.0938 | -82.3297
4.0941 | -82.3321
4.0945 | -82.3334
4.0949 | -82.3354
4.0952 | -82.3362
4.0956 | -82.3393
4.096 | -82.3407
4.0963 | -82.3427
4.0967 | -82.3473
4.097 | -82.3494
4.0974 | -82.3498
4.0978 | -82.3511
4.0981 | -82.3541
4.0985 | -82.3556
4.0988 | -82.3554
4.0992 | -82.3555
4.0996 | -82.3553
4.0999 | -82.357
4.1003 | -82.3587
4.1007 | -82.3595
4.101 | -82.36
4.1014 | -82.3607
4.1017 | -82.3619
4.1021 | -82.3644
4.1025 | -82.3662
4.1028 | -82.3687
4.1032 | -82.3703
4.1035 | -82.3711
4.1039 | -82.3733
4.1043 | -82.3743
4.1046 | -82.3735
4.105 | -82.373
4.1054 | -82.3749
4.1057 | -82.376
4.1061 | -82.3778
4.1064 | -82.3788
4.1068 | -82.38
4.1072 | -82.3799
4.1075 | -82.3804
4.1079 | -82.3822
4.1083 | -82.3831
4.1086 | -82.3838
4.109 | -82.3846
4.1093 | -82.3854
4.1097 | -82.3857
4.1101 | -82.3873
4.1104 | -82.3878
4.1108 | -82.3868
4.1111 | -82.3874
4.1115 | -82.3883
4.1119 | -82.3883
4.1122 | -82.3886
4.1126 | -82.3896
4.113 | -82.3914
4.1133 | -82.3922
4.1137 | -82.3929
4.114 | -82.3924
4.1144 | -82.3925
4.1148 | -82.3946
4.1151 | -82.399
4.1155 | -82.4004
4.1158 | -82.399
4.1162 | -82.3984
4.1166 | -82.4
4.1169 | -82.4006
4.1173 | -82.3999
4.1177 | -82.4004
4.118 | -82.4014
4.1184 | -82.4026
4.1187 | -82.4035
4.1191 | -82.4047
4.1195 | -82.4074
4.1198 | -82.4085
4.1202 | -82.4094
4.1206 | -82.4101
4.1209 | -82.409
4.1213 | -82.4073
4.1216 | -82.4067
4.122 | -82.407
4.1224 | -82.4092
4.1227 | -82.4112
4.1231 | -82.4118
4.1234 | -82.4139
4.1238 | -82.4142
4.1242 | -82.4151
4.1245 | -82.4158
4.1249 | -82.4166
4.1253 | -82.4177
4.1256 | -82.4184
4.126 | -82.4186
4.1263 | -82.4191
4.1267 | -82.4195
4.1271 | -82.4202
4.1274 | -82.4249
4.1278 | -82.4274
4.1281 | -82.428
4.1285 | -82.4282
4.1289 | -82.428
4.1292 | -82.4287
4.1296 | -82.4305
4.13 | -82.4332
4.1303 | -82.4329
4.1307 | -82.4333
4.131 | -82.4339
4.1314 | -82.4334
4.1318 | -82.4339
4.1321 | -82.4348
4.1325 | -82.4349
4.1329 | -82.4363
4.1332 | -82.4372
4.1336 | -82.4353
4.1339 | -82.434
4.1343 | -82.4316
4.1347 | -82.431
4.135 | -82.4316
4.1354 | -82.4316
4.1357 | -82.4301
4.1361 | -82.4305
4.1365 | -82.4295
4.1368 | -82.4289
4.1372 | -82.4293
4.1376 | -82.4283
4.1379 | -82.4274
4.1383 | -82.4272
4.1386 | -82.4273
4.139 | -82.4276
4.1394 | -82.427
4.1397 | -82.4268
4.1401 | -82.4266
4.1405 | -82.4269
4.1408 | -82.4273
4.1412 | -82.4271
4.1415 | -82.427
4.1419 | -82.4269
4.1423 | -82.4257
4.1426 | -82.4264
4.143 | -82.4264
4.1433 | -82.4265
4.1437 | -82.4264
4.1441 | -82.424
4.1444 | -82.4253
4.1448 | -82.4273
4.1452 | -82.4271
4.1455 | -82.4272
4.1459 | -82.4282
4.1462 | -82.4307
4.1466 | -82.4318
4.147 | -82.4318
4.1473 | -82.4315
4.1477 | -82.4322
4.148 | -82.4338
4.1484 | -82.4351
4.1488 | -82.437
4.1491 | -82.4395
4.1495 | -82.4392
4.1499 | -82.4379
4.1502 | -82.4379
4.1506 | -82.4375
4.1509 | -82.4372
4.1513 | -82.4377
4.1517 | -82.4386
4.152 | -82.4393
4.1524 | -82.4391
4.1528 | -82.4382
4.1531 | -82.4373
4.1535 | -82.4369
4.1538 | -82.4357
4.1542 | -82.4351
4.1546 | -82.4346
4.1549 | -82.4332
4.1553 | -82.4313
4.1556 | -82.4301
4.156 | -82.4294
4.1564 | -82.4284
4.1567 | -82.4274
4.1571 | -82.4273
4.1575 | -82.4268
4.1578 | -82.4265
4.1582 | -82.4263
4.1585 | -82.4265
4.1589 | -82.4262
4.1593 | -82.4262
4.1596 | -82.426
4.16 | -82.4261
4.1604 | -82.4257
4.1607 | -82.4266
4.1611 | -82.426
4.1614 | -82.4249
4.1618 | -82.4255
4.1622 | -82.4246
4.1625 | -82.4214
4.1629 | -82.4196
4.1632 | -82.4191
4.1636 | -82.4178
4.164 | -82.4172
4.1643 | -82.4174
4.1647 | -82.4189
4.1651 | -82.4177
4.1654 | -82.4174
4.1658 | -82.4165
4.1661 | -82.4169
4.1665 | -82.4186
4.1669 | -82.4179
4.1672 | -82.4174
4.1676 | -82.4171
4.1679 | -82.4172
4.1683 | -82.4168
4.1687 | -82.4167
4.169 | -82.4163
4.1694 | -82.4158
4.1698 | -82.4152
4.1701 | -82.4142
4.1705 | -82.4128
4.1708 | -82.4116
4.1712 | -82.4095
4.1716 | -82.4075
4.1719 | -82.4065
4.1723 | -82.4052
4.1727 | -82.4056
4.173 | -82.4084
4.1734 | -82.41
4.1737 | -82.4091
4.1741 | -82.4087
4.1745 | -82.4079
4.1748 | -82.4069
4.1752 | -82.4054
4.1755 | -82.4044
4.1759 | -82.4034
4.1763 | -82.4028
4.1766 | -82.4027
4.177 | -82.4015
4.1774 | -82.4012
4.1777 | -82.4004
4.1781 | -82.3996
4.1784 | -82.3988
4.1788 | -82.3987
4.1792 | -82.399
4.1795 | -82.3976
4.1799 | -82.397
4.1802 | -82.3965
4.1806 | -82.3963
4.181 | -82.397
4.1813 | -82.3983
4.1817 | -82.3996
4.1821 | -82.3993
4.1824 | -82.3976
4.1828 | -82.3951
4.1831 | -82.3969
4.1835 | -82.3955
4.1839 | -82.393
4.1842 | -82.3911
4.1846 | -82.3903
4.185 | -82.3898
4.1853 | -82.3887
4.1857 | -82.3879
4.186 | -82.387
4.1864 | -82.386
4.1868 | -82.3859
4.1871 | -82.3865
4.1875 | -82.3861
4.1878 | -82.384
4.1882 | -82.3809
4.1886 | -82.3789
4.1889 | -82.3772
4.1893 | -82.3753
4.1897 | -82.3739
4.19 | -82.3735
4.1904 | -82.3721
4.1907 | -82.3703
4.1911 | -82.3681
4.1915 | -82.3663
4.1918 | -82.3647
4.1922 | -82.364
4.1926 | -82.3637
4.1929 | -82.3626
4.1933 | -82.3616
4.1936 | -82.3603
4.194 | -82.3577
4.1944 | -82.3546
4.1947 | -82.3534
4.1951 | -82.3521
4.1954 | -82.3514
4.1958 | -82.3504
4.1962 | -82.3491
4.1965 | -82.3481
4.1969 | -82.3477
4.1973 | -82.3477
4.1976 | -82.3468
4.198 | -82.3456
4.1983 | -82.3446
4.1987 | -82.3444
4.1991 | -82.3436
4.1994 | -82.3414
4.1998 | -82.3432
4.2001 | -82.3429
4.2005 | -82.3415
4.2009 | -82.3402
4.2012 | -82.3393
4.2016 | -82.339
4.202 | -82.3386
4.2023 | -82.3379
4.2027 | -82.3342
4.203 | -82.3313
4.2034 | -82.3304
4.2038 | -82.3292
4.2041 | -82.3269
4.2045 | -82.3245
4.2049 | -82.3202
4.2052 | -82.3189
4.2056 | -82.3179
4.2059 | -82.3162
4.2063 | -82.3144
4.2067 | -82.3131
4.207 | -82.3117
4.2074 | -82.3119
4.2077 | -82.3134
4.2081 | -82.313
4.2085 | -82.3119
4.2088 | -82.3115
4.2092 | -82.3095
4.2096 | -82.3074
4.2099 | -82.3058
4.2103 | -82.3032
4.2106 | -82.3026
4.211 | -82.3013
4.2114 | -82.2996
4.2117 | -82.298
4.2121 | -82.2969
4.2124 | -82.296
4.2128 | -82.2951
4.2132 | -82.295
4.2135 | -82.2947
4.2139 | -82.293
4.2143 | -82.2924
4.2146 | -82.2916
4.215 | -82.2903
4.2153 | -82.2871
4.2157 | -82.2843
4.2161 | -82.2827
4.2164 | -82.2817
4.2168 | -82.2797
4.2172 | -82.2785
4.2175 | -82.2778
4.2179 | -82.2768
4.2182 | -82.2759
4.2186 | -82.2749
4.219 | -82.2736
4.2193 | -82.2719
4.2197 | -82.27
4.22 | -82.2683
4.2204 | -82.2664
4.2208 | -82.265
4.2211 | -82.2639
4.2215 | -82.263
4.2219 | -82.2619
4.2222 | -82.2612
4.2226 | -82.2599
4.2229 | -82.2591
4.2233 | -82.2589
4.2237 | -82.2572
4.224 | -82.2549
4.2244 | -82.254
4.2248 | -82.2527
4.2251 | -82.2516
4.2255 | -82.2507
4.2258 | -82.2504
4.2262 | -82.2497
4.2266 | -82.249
4.2269 | -82.2476
4.2273 | -82.2464
4.2276 | -82.2439
4.228 | -82.2422
4.2284 | -82.2419
4.2287 | -82.2413
4.2291 | -82.2399
4.2295 | -82.2376
4.2298 | -82.2358
4.2302 | -82.2354
4.2305 | -82.2341
4.2309 | -82.2329
4.2313 | -82.232
4.2316 | -82.2315
4.232 | -82.2309
4.2323 | -82.2303
4.2327 | -82.2291
4.2331 | -82.2253
4.2334 | -82.2228
4.2338 | -82.2223
4.2342 | -82.2229
4.2345 | -82.2226
4.2349 | -82.2207
4.2352 | -82.2199
4.2356 | -82.2212
4.236 | -82.2204
4.2363 | -82.2202
4.2367 | -82.219
4.2371 | -82.218
4.2374 | -82.217
4.2378 | -82.2157
4.2381 | -82.2158
4.2385 | -82.2147
4.2389 | -82.2141
4.2392 | -82.2134
4.2396 | -82.2122
4.2399 | -82.2107
4.2403 | -82.2096
4.2407 | -82.2095
4.241 | -82.2091
4.2414 | -82.2088
4.2418 | -82.2083
4.2421 | -82.2076
4.2425 | -82.2081
4.2428 | -82.2085
4.2432 | -82.2081
4.2436 | -82.2086
4.2439 | -82.2081
4.2443 | -82.2076
4.2447 | -82.2068
4.245 | -82.2059
4.2454 | -82.2052
4.2457 | -82.2042
4.2461 | -82.2041
4.2465 | -82.2035
4.2468 | -82.2033
4.2472 | -82.203
4.2475 | -82.2018
4.2479 | -82.2013
4.2483 | -82.2006
4.2486 | -82.1998
4.249 | -82.2009
4.2494 | -82.2015
4.2497 | -82.2016
4.2501 | -82.2014
4.2504 | -82.2016
4.2508 | -82.2032
4.2512 | -82.2047
4.2515 | -82.2053
4.2519 | -82.2056
4.2522 | -82.2069
4.2526 | -82.2065
4.253 | -82.2057
4.2533 | -82.2043
4.2537 | -82.2037
4.2541 | -82.2041
4.2544 | -82.2035
4.2548 | -82.203
4.2551 | -82.203
4.2555 | -82.2043
4.2559 | -82.2049
4.2562 | -82.2038
4.2566 | -82.2029
4.257 | -82.2014
4.2573 | -82.2004
4.2577 | -82.2003
4.258 | -82.2007
4.2584 | -82.2008
4.2588 | -82.1995
4.2591 | -82.1994
4.2595 | -82.1997
4.2598 | -82.1976
4.2602 | -82.1964
4.2606 | -82.1955
4.2609 | -82.1946
4.2613 | -82.194
4.2617 | -82.1943
4.262 | -82.195
4.2624 | -82.1947
4.2627 | -82.1945
4.2631 | -82.1938
4.2635 | -82.1929
4.2638 | -82.1923
4.2642 | -82.1915
4.2645 | -82.1903
4.2649 | -82.1894
4.2653 | -82.1889
4.2656 | -82.1885
4.266 | -82.188
4.2664 | -82.1872
4.2667 | -82.1863
4.2671 | -82.185
4.2674 | -82.183
4.2678 | -82.1812
4.2682 | -82.1802
4.2685 | -82.1801
4.2689 | -82.18
4.2693 | -82.1811
4.2696 | -82.1804
4.27 | -82.1794
4.2703 | -82.1792
4.2707 | -82.1791
4.2711 | -82.1788
4.2714 | -82.1782
4.2718 | -82.1782
4.2721 | -82.1773
4.2725 | -82.1764
4.2729 | -82.1762
4.2732 | -82.1758
4.2736 | -82.1748
4.274 | -82.1732
4.2743 | -82.1728
4.2747 | -82.173
4.275 | -82.1738
4.2754 | -82.1726
4.2758 | -82.1715
4.2761 | -82.1712
4.2765 | -82.1709
4.2768 | -82.1709
4.2772 | -82.171
4.2776 | -82.1718
4.2779 | -82.1729
4.2783 | -82.1732
4.2787 | -82.1738
4.279 | -82.1755
4.2794 | -82.1764
4.2797 | -82.1772
4.2801 | -82.1775
4.2805 | -82.1774
4.2808 | -82.1771
4.2812 | -82.1767
4.2816 | -82.1772
4.2819 | -82.1773
4.2823 | -82.1766
4.2826 | -82.1766
4.283 | -82.1772
4.2834 | -82.1778
4.2837 | -82.1782
4.2841 | -82.1782
4.2844 | -82.1786
4.2848 | -82.1787
4.2852 | -82.1779
4.2855 | -82.1762
4.2859 | -82.1755
4.2863 | -82.1757
4.2866 | -82.1756
4.287 | -82.1754
4.2873 | -82.1753
4.2877 | -82.1756
4.2881 | -82.1766
4.2884 | -82.1765
4.2888 | -82.1763
4.2892 | -82.1763
4.2895 | -82.1769
4.2899 | -82.1775
4.2902 | -82.1779
4.2906 | -82.1778
4.291 | -82.1773
4.2913 | -82.1768
4.2917 | -82.1756
4.292 | -82.1745
4.2924 | -82.1756
4.2928 | -82.1759
4.2931 | -82.1753
4.2935 | -82.1754
4.2939 | -82.1752
4.2942 | -82.1745
4.2946 | -82.1761
4.2949 | -82.1765
4.2953 | -82.1762
4.2957 | -82.1764
4.296 | -82.1764
4.2964 | -82.1747
4.2967 | -82.1744
4.2971 | -82.1746
4.2975 | -82.1746
4.2978 | -82.1731
4.2982 | -82.1723
4.2986 | -82.1724
4.2989 | -82.173
4.2993 | -82.1731
4.2996 | -82.1734
4.3 | -82.1745
4.3004 | -82.1747
4.3007 | -82.1759
4.3011 | -82.1753
4.3015 | -82.1749
4.3018 | -82.1746
4.3022 | -82.1745
4.3025 | -82.1734
4.3029 | -82.1735
4.3033 | -82.1735
4.3036 | -82.1734
4.304 | -82.1734
4.3043 | -82.1732
4.3047 | -82.173
4.3051 | -82.1728
4.3054 | -82.172
4.3058 | -82.1711
4.3062 | -82.1692
4.3065 | -82.1686
4.3069 | -82.1688
4.3072 | -82.1685
4.3076 | -82.1686
4.308 | -82.1689
4.3083 | -82.1679
4.3087 | -82.1663
4.3091 | -82.1664
4.3094 | -82.1666
4.3098 | -82.1671
4.3101 | -82.1676
4.3105 | -82.1671
4.3109 | -82.1672
4.3112 | -82.1659
4.3116 | -82.1664
4.3119 | -82.1669
4.3123 | -82.1675
4.3127 | -82.1674
4.313 | -82.1671
4.3134 | -82.1679
4.3138 | -82.1687
4.3141 | -82.1684
4.3145 | -82.1683
4.3148 | -82.1685
4.3152 | -82.1683
4.3156 | -82.1683
4.3159 | -82.1675
4.3163 | -82.1664
4.3166 | -82.1665
4.317 | -82.1667
4.3174 | -82.1673
4.3177 | -82.1678
4.3181 | -82.1683
4.3185 | -82.1685
4.3188 | -82.1679
4.3192 | -82.1682
4.3195 | -82.1685
4.3199 | -82.1686
4.3203 | -82.1689
4.3206 | -82.1702
4.321 | -82.1717
4.3214 | -82.1712
4.3217 | -82.1713
4.3221 | -82.1717
4.3224 | -82.1712
4.3228 | -82.1724
4.3232 | -82.1728
4.3235 | -82.172
4.3239 | -82.1719
4.3242 | -82.1718
4.3246 | -82.1721
4.325 | -82.1726
4.3253 | -82.1729
4.3257 | -82.1736
4.3261 | -82.1746
4.3264 | -82.1757
4.3268 | -82.1765
4.3271 | -82.1764
4.3275 | -82.1772
4.3279 | -82.1782
4.3282 | -82.1788
4.3286 | -82.178
4.3289 | -82.1769
4.3293 | -82.1757
4.3297 | -82.1775
4.33 | -82.179
4.3304 | -82.1815
4.3308 | -82.1851
4.3311 | -82.1852
4.3315 | -82.1854
4.3318 | -82.1833
4.3322 | -82.1839
4.3326 | -82.1843
4.3329 | -82.1853
4.3333 | -82.1864
4.3337 | -82.1867
4.334 | -82.1872
4.3344 | -82.1864
4.3347 | -82.1848
4.3351 | -82.1846
4.3355 | -82.185
4.3358 | -82.1857
4.3362 | -82.1856
4.3365 | -82.1859
4.3369 | -82.1873
4.3373 | -82.1891
4.3376 | -82.1888
4.338 | -82.187
4.3384 | -82.1868
4.3387 | -82.187
4.3391 | -82.1874
4.3394 | -82.1887
4.3398 | -82.189
4.3402 | -82.1891
4.3405 | -82.1886
4.3409 | -82.187
4.3413 | -82.1858
4.3416 | -82.1845
4.342 | -82.1828
4.3423 | -82.1817
4.3427 | -82.1819
4.3431 | -82.1822
4.3434 | -82.1825
4.3438 | -82.1837
4.3441 | -82.185
4.3445 | -82.1857
4.3449 | -82.1868
4.3452 | -82.1877
4.3456 | -82.1883
4.346 | -82.1894
4.3463 | -82.1903
4.3467 | -82.1909
4.347 | -82.1912
4.3474 | -82.1919
4.3478 | -82.1929
4.3481 | -82.1938
4.3485 | -82.194
4.3488 | -82.1946
4.3492 | -82.1947
4.3496 | -82.1951
4.3499 | -82.1962
4.3503 | -82.1974
4.3507 | -82.1985
4.351 | -82.1983
4.3514 | -82.1987
4.3517 | -82.1977
4.3521 | -82.1961
4.3525 | -82.1962
4.3528 | -82.1963
4.3532 | -82.197
4.3536 | -82.1973
4.3539 | -82.1978
4.3543 | -82.1991
4.3546 | -82.1998
4.355 | -82.2
4.3554 | -82.2009
4.3557 | -82.2015
4.3561 | -82.2006
4.3564 | -82.2007
4.3568 | -82.2019
4.3572 | -82.2029
4.3575 | -82.2031
4.3579 | -82.2042
4.3583 | -82.205
4.3586 | -82.2051
4.359 | -82.2054
4.3593 | -82.2058
4.3597 | -82.2063
4.3601 | -82.207
4.3604 | -82.2078
4.3608 | -82.2081
4.3612 | -82.209
4.3615 | -82.2087
4.3619 | -82.2078
4.3622 | -82.2075
4.3626 | -82.2085
4.363 | -82.2094
4.3633 | -82.2109
4.3637 | -82.2116
4.364 | -82.212
4.3644 | -82.2117
4.3648 | -82.2097
4.3651 | -82.2098
4.3655 | -82.2107
4.3659 | -82.2116
4.3662 | -82.2124
4.3666 | -82.2126
4.3669 | -82.2126
4.3673 | -82.2138
4.3677 | -82.2142
4.368 | -82.2134
4.3684 | -82.2116
4.3687 | -82.2116
4.3691 | -82.2116
4.3695 | -82.2119
4.3698 | -82.2121
4.3702 | -82.2116
4.3706 | -82.2122
4.3709 | -82.2124
4.3713 | -82.213
4.3716 | -82.213
4.372 | -82.2136
4.3724 | -82.2143
4.3727 | -82.216
4.3731 | -82.2162
4.3735 | -82.2163
4.3738 | -82.2172
4.3742 | -82.2174
4.3745 | -82.2179
4.3749 | -82.218
4.3753 | -82.2181
4.3756 | -82.219
4.376 | -82.22
4.3763 | -82.2207
4.3767 | -82.2215
4.3771 | -82.2222
4.3774 | -82.2235
4.3778 | -82.2232
4.3782 | -82.2238
4.3785 | -82.2243
4.3789 | -82.2244
4.3792 | -82.2233
4.3796 | -82.2216
4.38 | -82.2208
4.3803 | -82.2212
4.3807 | -82.2218
4.381 | -82.221
4.3814 | -82.2213
4.3818 | -82.2217
4.3821 | -82.2219
4.3825 | -82.2224
4.3829 | -82.2227
4.3832 | -82.2228
4.3836 | -82.2233
4.3839 | -82.2236
4.3843 | -82.2219
4.3847 | -82.2218
4.385 | -82.2225
4.3854 | -82.2242
4.3858 | -82.2264
4.3861 | -82.2255
4.3865 | -82.2254
4.3868 | -82.2259
4.3872 | -82.2257
4.3876 | -82.2265
4.3879 | -82.227
4.3883 | -82.2271
4.3886 | -82.2271
4.389 | -82.2273
4.3894 | -82.2277
4.3897 | -82.2281
4.3901 | -82.2294
4.3905 | -82.23
4.3908 | -82.2308
4.3912 | -82.2315
4.3915 | -82.2319
4.3919 | -82.2325
4.3923 | -82.2338
4.3926 | -82.2344
4.393 | -82.2346
4.3933 | -82.2359
4.3937 | -82.2379
4.3941 | -82.2398
4.3944 | -82.24
4.3948 | -82.241
4.3952 | -82.2402
4.3955 | -82.24
4.3959 | -82.2403
4.3962 | -82.2414
4.3966 | -82.2424
4.397 | -82.2431
4.3973 | -82.2435
4.3977 | -82.244
4.3981 | -82.2443
4.3984 | -82.2449
4.3988 | -82.2452
4.3991 | -82.2465
4.3995 | -82.2464
4.3999 | -82.2464
4.4002 | -82.2482
4.4006 | -82.2505
4.4009 | -82.2507
4.4013 | -82.2501
4.4017 | -82.2496
4.402 | -82.2506
4.4024 | -82.2513
4.4028 | -82.2529
4.4031 | -82.2545
4.4035 | -82.2547
4.4038 | -82.2537
4.4042 | -82.2545
4.4046 | -82.2548
4.4049 | -82.2546
4.4053 | -82.254
4.4057 | -82.2544
4.406 | -82.2542
4.4064 | -82.2555
4.4067 | -82.2581
4.4071 | -82.2589
4.4075 | -82.26
4.4078 | -82.2616
4.4082 | -82.2605
4.4085 | -82.2576
4.4089 | -82.2571
4.4093 | -82.2563
4.4096 | -82.2567
4.41 | -82.2579
4.4104 | -82.2595
4.4107 | -82.2606
4.4111 | -82.261
4.4114 | -82.2611
4.4118 | -82.2617
4.4122 | -82.2634
4.4125 | -82.2638
4.4129 | -82.2634
4.4132 | -82.2639
4.4136 | -82.2657
4.414 | -82.2662
4.4143 | -82.2665
4.4147 | -82.2656
4.4151 | -82.2655
4.4154 | -82.2653
4.4158 | -82.2649
4.4161 | -82.2651
4.4165 | -82.2653
4.4169 | -82.265
4.4172 | -82.2653
4.4176 | -82.2672
4.418 | -82.2694
4.4183 | -82.2708
4.4187 | -82.2716
4.419 | -82.2687
4.4194 | -82.2668
4.4198 | -82.2661
4.4201 | -82.2663
4.4205 | -82.2668
4.4208 | -82.2676
4.4212 | -82.2679
4.4216 | -82.2682
4.4219 | -82.2685
4.4223 | -82.2679
4.4227 | -82.2683
4.423 | -82.2686
4.4234 | -82.2689
4.4237 | -82.2696
4.4241 | -82.2713
4.4245 | -82.2718
4.4248 | -82.2722
4.4252 | -82.2732
4.4256 | -82.2753
4.4259 | -82.2768
4.4263 | -82.2786
4.4266 | -82.279
4.427 | -82.2784
4.4274 | -82.2771
4.4277 | -82.2775
4.4281 | -82.2757
4.4284 | -82.2729
4.4288 | -82.2737
4.4292 | -82.2743
4.4295 | -82.2747
4.4299 | -82.2749
4.4303 | -82.2754
4.4306 | -82.276
4.431 | -82.2774
4.4313 | -82.2789
4.4317 | -82.2805
4.4321 | -82.2831
4.4324 | -82.2842
4.4328 | -82.2836
4.4331 | -82.2832
4.4335 | -82.2849
4.4339 | -82.2847
4.4342 | -82.2856
4.4346 | -82.2859
4.435 | -82.2861
4.4353 | -82.2864
4.4357 | -82.2859
4.436 | -82.2861
4.4364 | -82.2865
4.4368 | -82.2859
4.4371 | -82.284
4.4375 | -82.2837
4.4379 | -82.2829
4.4382 | -82.2827
4.4386 | -82.2827
4.4389 | -82.2823
4.4393 | -82.2824
4.4397 | -82.2829
4.44 | -82.2832
4.4404 | -82.281
4.4407 | -82.2762
4.4411 | -82.2741
4.4415 | -82.2743
4.4418 | -82.2752
4.4422 | -82.2769
4.4426 | -82.2779
4.4429 | -82.2781
4.4433 | -82.2783
4.4436 | -82.2787
4.444 | -82.2787
4.4444 | -82.2787
4.4447 | -82.2793
4.4451 | -82.2799
4.4454 | -82.2806
4.4458 | -82.2814
4.4462 | -82.2816
4.4465 | -82.2819
4.4469 | -82.283
4.4473 | -82.2826
4.4476 | -82.2826
4.448 | -82.2833
4.4483 | -82.285
4.4487 | -82.2857
4.4491 | -82.2859
4.4494 | -82.2865
4.4498 | -82.2872
4.4502 | -82.2874
4.4505 | -82.2878
4.4509 | -82.2881
4.4512 | -82.2888
4.4516 | -82.289
4.452 | -82.2898
4.4523 | -82.2897
4.4527 | -82.2906
4.453 | -82.2912
4.4534 | -82.2915
4.4538 | -82.2923
4.4541 | -82.2934
4.4545 | -82.2949
4.4549 | -82.2957
4.4552 | -82.2957
4.4556 | -82.2966
4.4559 | -82.2972
4.4563 | -82.2957
4.4567 | -82.2952
4.457 | -82.2958
4.4574 | -82.2949
4.4578 | -82.295
4.4581 | -82.2951
4.4585 | -82.2949
4.4588 | -82.295
4.4592 | -82.293
4.4596 | -82.2923
4.4599 | -82.2925
4.4603 | -82.2917
4.4606 | -82.292
4.461 | -82.293
4.4614 | -82.2942
4.4617 | -82.2961
4.4621 | -82.2998
4.4625 | -82.3012
4.4628 | -82.3014
4.4632 | -82.3021
4.4635 | -82.3023
4.4639 | -82.3026
4.4643 | -82.3029
4.4646 | -82.3034
4.465 | -82.3033
4.4653 | -82.3024
4.4657 | -82.3021
4.4661 | -82.3016
4.4664 | -82.3014
4.4668 | -82.3013
4.4672 | -82.3003
4.4675 | -82.2994
4.4679 | -82.2982
4.4682 | -82.2974
4.4686 | -82.2975
4.469 | -82.2978
4.4693 | -82.2978
4.4697 | -82.2979
4.4701 | -82.2982
4.4704 | -82.299
4.4708 | -82.2998
4.4711 | -82.3006
4.4715 | -82.3018
4.4719 | -82.3033
4.4722 | -82.3034
4.4726 | -82.3027
4.4729 | -82.3021
4.4733 | -82.3023
4.4737 | -82.3029
4.474 | -82.3032
4.4744 | -82.3034
4.4748 | -82.3042
4.4751 | -82.3052
4.4755 | -82.3058
4.4758 | -82.3071
4.4762 | -82.3085
4.4766 | -82.3096
4.4769 | -82.3103
4.4773 | -82.3109
4.4776 | -82.3119
4.478 | -82.313
4.4784 | -82.3134
4.4787 | -82.3136
4.4791 | -82.3132
4.4795 | -82.3135
4.4798 | -82.314
4.4802 | -82.3144
4.4805 | -82.3151
4.4809 | -82.3143
4.4813 | -82.3122
4.4816 | -82.3095
4.482 | -82.3092
4.4824 | -82.3092
4.4827 | -82.3093
4.4831 | -82.3094
4.4834 | -82.3105
4.4838 | -82.3127
4.4842 | -82.3122
4.4845 | -82.3138
4.4849 | -82.3149
4.4852 | -82.3146
4.4856 | -82.3139
4.486 | -82.3142
4.4863 | -82.3149
4.4867 | -82.3149
4.4871 | -82.3147
4.4874 | -82.3145
4.4878 | -82.3146
4.4881 | -82.3149
4.4885 | -82.3159
4.4889 | -82.3161
4.4892 | -82.3169
4.4896 | -82.3178
4.49 | -82.3199
4.4903 | -82.3217
4.4907 | -82.3232
4.491 | -82.3243
4.4914 | -82.3249
4.4918 | -82.3244
4.4921 | -82.3243
4.4925 | -82.3253
4.4928 | -82.3263
4.4932 | -82.3264
4.4936 | -82.3259
4.4939 | -82.3261
4.4943 | -82.3261
4.4947 | -82.3265
4.495 | -82.3272
4.4954 | -82.3276
4.4957 | -82.3293
4.4961 | -82.3306
4.4965 | -82.3311
4.4968 | -82.3308
4.4972 | -82.3303
4.4975 | -82.3304
4.4979 | -82.3312
4.4983 | -82.3321
4.4986 | -82.3323
4.499 | -82.3333
4.4994 | -82.3334
4.4997 | -82.3337
4.5001 | -82.334
4.5004 | -82.3345
4.5008 | -82.3344
4.5012 | -82.3345
4.5015 | -82.3344
4.5019 | -82.3354
4.5023 | -82.3351
4.5026 | -82.3361
4.503 | -82.3363
4.5033 | -82.3367
4.5037 | -82.337
4.5041 | -82.3378
4.5044 | -82.3383
4.5048 | -82.3394
4.5051 | -82.3423
4.5055 | -82.3445
4.5059 | -82.3446
4.5062 | -82.3448
4.5066 | -82.3461
4.507 | -82.3473
4.5073 | -82.3507
4.5077 | -82.3514
4.508 | -82.3511
4.5084 | -82.3503
4.5088 | -82.3507
4.5091 | -82.3517
4.5095 | -82.3516
4.5099 | -82.3521
4.5102 | -82.3526
4.5106 | -82.3525
4.5109 | -82.3529
4.5113 | -82.3528
4.5117 | -82.3533
4.512 | -82.3524
4.5124 | -82.3518
4.5127 | -82.3529
4.5131 | -82.3535
4.5135 | -82.3539
4.5138 | -82.3541
4.5142 | -82.3544
4.5146 | -82.355
4.5149 | -82.3554
4.5153 | -82.3557
4.5156 | -82.3559
4.516 | -82.3558
4.5164 | -82.3562
4.5167 | -82.3563
4.5171 | -82.3554
4.5174 | -82.3532
4.5178 | -82.3527
4.5182 | -82.3521
4.5185 | -82.3536
4.5189 | -82.3558
4.5193 | -82.3566
4.5196 | -82.3566
4.52 | -82.3563
4.5203 | -82.3564
4.5207 | -82.3563
4.5211 | -82.3557
4.5214 | -82.3562
4.5218 | -82.3559
4.5222 | -82.3558
4.5225 | -82.3558
4.5229 | -82.3564
4.5232 | -82.3574
4.5236 | -82.3584
4.524 | -82.3593
4.5243 | -82.3596
4.5247 | -82.3597
4.525 | -82.3592
4.5254 | -82.359
4.5258 | -82.3594
4.5261 | -82.3599
4.5265 | -82.3606
4.5269 | -82.3612
4.5272 | -82.3611
4.5276 | -82.3603
4.5279 | -82.361
4.5283 | -82.3616
4.5287 | -82.362
4.529 | -82.3622
4.5294 | -82.3626
4.5297 | -82.3625
4.5301 | -82.3629
4.5305 | -82.3631
4.5308 | -82.3628
4.5312 | -82.3626
4.5316 | -82.3626
4.5319 | -82.3633
4.5323 | -82.3655
4.5326 | -82.3667
4.533 | -82.3676
4.5334 | -82.368
4.5337 | -82.3686
4.5341 | -82.3705
4.5345 | -82.3717
4.5348 | -82.3731
4.5352 | -82.3727
4.5355 | -82.3719
4.5359 | -82.3718
4.5363 | -82.372
4.5366 | -82.3725
4.537 | -82.373
4.5373 | -82.3732
4.5377 | -82.3761
4.5381 | -82.3775
4.5384 | -82.3764
4.5388 | -82.3751
4.5392 | -82.375
4.5395 | -82.3763
4.5399 | -82.3774
4.5402 | -82.3772
4.5406 | -82.3779
4.541 | -82.3779
4.5413 | -82.3778
4.5417 | -82.3785
4.542 | -82.3791
4.5424 | -82.3796
4.5428 | -82.3803
4.5431 | -82.3801
4.5435 | -82.3789
4.5439 | -82.3789
4.5442 | -82.3789
4.5446 | -82.3792
4.5449 | -82.3788
4.5453 | -82.3785
4.5457 | -82.379
4.546 | -82.3789
4.5464 | -82.3778
4.5468 | -82.3769
4.5471 | -82.377
4.5475 | -82.3773
4.5478 | -82.3777
4.5482 | -82.3772
4.5486 | -82.3771
4.5489 | -82.3777
4.5493 | -82.3778
4.5496 | -82.3767
4.55 | -82.3773
4.5504 | -82.378
4.5507 | -82.3787
4.5511 | -82.38
4.5515 | -82.3812
4.5518 | -82.3815
4.5522 | -82.3816
4.5525 | -82.3816
4.5529 | -82.3819
4.5533 | -82.3831
4.5536 | -82.3841
4.554 | -82.3846
4.5544 | -82.3857
4.5547 | -82.3867
4.5551 | -82.3871
4.5554 | -82.3875
4.5558 | -82.3873
4.5562 | -82.3877
4.5565 | -82.3882
4.5569 | -82.3886
4.5572 | -82.3889
4.5576 | -82.3895
4.558 | -82.3894
4.5583 | -82.3878
4.5587 | -82.3882
4.5591 | -82.3887
4.5594 | -82.3893
4.5598 | -82.3903
4.5601 | -82.3912
4.5605 | -82.3919
4.5609 | -82.3926
4.5612 | -82.3929
4.5616 | -82.3942
4.5619 | -82.3945
4.5623 | -82.3951
4.5627 | -82.3963
4.563 | -82.397
4.5634 | -82.3986
4.5638 | -82.3997
4.5641 | -82.3996
4.5645 | -82.4006
4.5648 | -82.4019
4.5652 | -82.4031
4.5656 | -82.4038
4.5659 | -82.4039
4.5663 | -82.4037
4.5667 | -82.4059
4.567 | -82.4077
4.5674 | -82.4074
4.5677 | -82.4079
4.5681 | -82.4087
4.5685 | -82.4097
4.5688 | -82.4102
4.5692 | -82.4102
4.5695 | -82.4104
4.5699 | -82.4111
4.5703 | -82.4117
4.5706 | -82.4126
4.571 | -82.4134
4.5714 | -82.4143
4.5717 | -82.4153
4.5721 | -82.4167
4.5724 | -82.4173
4.5728 | -82.4176
4.5732 | -82.4181
4.5735 | -82.4188
4.5739 | -82.42
4.5743 | -82.4218
4.5746 | -82.4241
4.575 | -82.426
4.5753 | -82.427
4.5757 | -82.4275
4.5761 | -82.428
4.5764 | -82.4285
4.5768 | -82.4294
4.5771 | -82.4308
4.5775 | -82.4324
4.5779 | -82.4332
4.5782 | -82.434
4.5786 | -82.4349
4.579 | -82.4361
4.5793 | -82.4375
4.5797 | -82.4385
4.58 | -82.4391
4.5804 | -82.4395
4.5808 | -82.4398
4.5811 | -82.4406
4.5815 | -82.4428
4.5818 | -82.4444
4.5822 | -82.4454
4.5826 | -82.4472
4.5829 | -82.4475
4.5833 | -82.4512
4.5837 | -82.4534
4.584 | -82.4543
4.5844 | -82.4553
4.5847 | -82.4564
4.5851 | -82.458
4.5855 | -82.4585
4.5858 | -82.4586
4.5862 | -82.4593
4.5866 | -82.4601
4.5869 | -82.4591
4.5873 | -82.4574
4.5876 | -82.4564
4.588 | -82.4576
4.5884 | -82.4578
4.5887 | -82.4564
4.5891 | -82.4559
4.5894 | -82.4568
4.5898 | -82.4566
4.5902 | -82.4578
4.5905 | -82.4597
4.5909 | -82.4601
4.5913 | -82.4599
4.5916 | -82.4605
4.592 | -82.4612
4.5923 | -82.4629
4.5927 | -82.4653
4.5931 | -82.4675
4.5934 | -82.4692
4.5938 | -82.4694
4.5941 | -82.4704
4.5945 | -82.4712
4.5949 | -82.4719
4.5952 | -82.4726
4.5956 | -82.4735
4.596 | -82.4736
4.5963 | -82.4738
4.5967 | -82.4747
4.597 | -82.4746
4.5974 | -82.4749
4.5978 | -82.4759
4.5981 | -82.4772
4.5985 | -82.4785
4.5989 | -82.4788
4.5992 | -82.48
4.5996 | -82.481
4.5999 | -82.4824
4.6003 | -82.4833
4.6007 | -82.4837
4.601 | -82.4851
4.6014 | -82.486
4.6017 | -82.4885
4.6021 | -82.4896
4.6025 | -82.489
4.6028 | -82.4877
4.6032 | -82.4837
4.6036 | -82.4838
4.6039 | -82.4844
4.6043 | -82.486
4.6046 | -82.4877
4.605 | -82.4882
4.6054 | -82.4877
4.6057 | -82.4899
4.6061 | -82.4901
4.6065 | -82.4895
4.6068 | -82.4901
4.6072 | -82.4915
4.6075 | -82.493
4.6079 | -82.4942
4.6083 | -82.4949
4.6086 | -82.4943
4.609 | -82.4947
4.6093 | -82.4958
4.6097 | -82.4969
4.6101 | -82.4967
4.6104 | -82.4974
4.6108 | -82.4988
4.6112 | -82.5004
4.6115 | -82.5019
4.6119 | -82.5025
4.6122 | -82.5039
4.6126 | -82.5052
4.613 | -82.5069
4.6133 | -82.5081
4.6137 | -82.5095
4.614 | -82.5094
4.6144 | -82.5098
4.6148 | -82.5108
4.6151 | -82.5121
4.6155 | -82.5132
4.6159 | -82.515
4.6162 | -82.518
4.6166 | -82.5193
4.6169 | -82.5202
4.6173 | -82.5216
4.6177 | -82.5226
4.618 | -82.5234
4.6184 | -82.5247
4.6188 | -82.525
4.6191 | -82.5244
4.6195 | -82.5247
4.6198 | -82.525
4.6202 | -82.525
4.6206 | -82.5257
4.6209 | -82.5265
4.6213 | -82.528
4.6216 | -82.5307
4.622 | -82.5316
4.6224 | -82.5327
4.6227 | -82.5333
4.6231 | -82.5346
4.6235 | -82.5359
4.6238 | -82.5364
4.6242 | -82.538
4.6245 | -82.539
4.6249 | -82.5408
4.6253 | -82.5407
4.6256 | -82.5396
4.626 | -82.5402
4.6264 | -82.5411
4.6267 | -82.5424
4.6271 | -82.544
4.6274 | -82.5448
4.6278 | -82.5453
4.6282 | -82.5458
4.6285 | -82.5465
4.6289 | -82.5472
4.6292 | -82.5475
4.6296 | -82.5493
4.63 | -82.552
4.6303 | -82.5531
4.6307 | -82.5535
4.6311 | -82.554
4.6314 | -82.5545
4.6318 | -82.5549
4.6321 | -82.556
4.6325 | -82.5582
4.6329 | -82.5591
4.6332 | -82.5602
4.6336 | -82.561
4.6339 | -82.562
4.6343 | -82.5638
4.6347 | -82.5649
4.635 | -82.5664
4.6354 | -82.567
4.6358 | -82.5681
4.6361 | -82.569
4.6365 | -82.5695
4.6368 | -82.5695
4.6372 | -82.5699
4.6376 | -82.5699
4.6379 | -82.5689
4.6383 | -82.5687
4.6387 | -82.5693
4.639 | -82.5698
4.6394 | -82.5705
4.6397 | -82.5709
4.6401 | -82.5704
4.6405 | -82.5717
4.6408 | -82.5743
4.6412 | -82.5748
4.6415 | -82.5758
4.6419 | -82.5756
4.6423 | -82.5759
4.6426 | -82.577
4.643 | -82.5773
4.6434 | -82.5779
4.6437 | -82.5806
4.6441 | -82.58
4.6444 | -82.5792
4.6448 | -82.579
4.6452 | -82.5776
4.6455 | -82.5784
4.6459 | -82.5791
4.6462 | -82.5797
4.6466 | -82.5812
4.647 | -82.5837
4.6473 | -82.5846
4.6477 | -82.5848
4.6481 | -82.5849
4.6484 | -82.586
4.6488 | -82.5861
4.6491 | -82.5866
4.6495 | -82.5871
4.6499 | -82.5881
4.6502 | -82.5905
4.6506 | -82.5913
4.651 | -82.5909
4.6513 | -82.5911
4.6517 | -82.5911
4.652 | -82.5914
4.6524 | -82.592
4.6528 | -82.5938
4.6531 | -82.5965
4.6535 | -82.5986
4.6538 | -82.6006
4.6542 | -82.6021
4.6546 | -82.6036
4.6549 | -82.6047
4.6553 | -82.6049
4.6557 | -82.6032
4.656 | -82.6035
4.6564 | -82.6037
4.6567 | -82.6037
4.6571 | -82.6035
4.6575 | -82.6023
4.6578 | -82.6009
4.6582 | -82.6005
4.6585 | -82.6005
4.6589 | -82.6009
4.6593 | -82.6009
4.6596 | -82.6012
4.66 | -82.6012
4.6604 | -82.5998
4.6607 | -82.5995
4.6611 | -82.5996
4.6614 | -82.6004
4.6618 | -82.5998
4.6622 | -82.5986
4.6625 | -82.599
4.6629 | -82.5995
4.6633 | -82.5991
4.6636 | -82.5984
4.664 | -82.5989
4.6643 | -82.6004
4.6647 | -82.6024
4.6651 | -82.6045
4.6654 | -82.6055
4.6658 | -82.6063
4.6661 | -82.6029
4.6665 | -82.6028
4.6669 | -82.6026
4.6672 | -82.6029
4.6676 | -82.6028
4.668 | -82.6028
4.6683 | -82.6033
4.6687 | -82.6035
4.669 | -82.6035
4.6694 | -82.6035
4.6698 | -82.6047
4.6701 | -82.6044
4.6705 | -82.603
4.6709 | -82.6029
4.6712 | -82.6029
4.6716 | -82.6028
4.6719 | -82.6029
4.6723 | -82.6027
4.6727 | -82.6028
4.673 | -82.6036
4.6734 | -82.604
4.6737 | -82.6041
4.6741 | -82.6046
4.6745 | -82.6039
4.6748 | -82.6039
4.6752 | -82.6047
4.6756 | -82.6056
4.6759 | -82.6061
4.6763 | -82.6059
4.6766 | -82.6047
4.677 | -82.6027
4.6774 | -82.6008
4.6777 | -82.6004
4.6781 | -82.5994
4.6784 | -82.5979
4.6788 | -82.598
4.6792 | -82.6003
4.6795 | -82.6013
4.6799 | -82.6015
4.6803 | -82.6016
4.6806 | -82.6011
4.681 | -82.6009
4.6813 | -82.601
4.6817 | -82.601
4.6821 | -82.6011
4.6824 | -82.6014
4.6828 | -82.6036
4.6832 | -82.6035
4.6835 | -82.6037
4.6839 | -82.6045
4.6842 | -82.6055
4.6846 | -82.6063
4.685 | -82.6072
4.6853 | -82.607
4.6857 | -82.6071
4.686 | -82.607
4.6864 | -82.6072
4.6868 | -82.6063
4.6871 | -82.6052
4.6875 | -82.6044
4.6879 | -82.6052
4.6882 | -82.6054
4.6886 | -82.6046
4.6889 | -82.6042
4.6893 | -82.6043
4.6897 | -82.6042
4.69 | -82.6036
4.6904 | -82.6035
4.6908 | -82.6031
4.6911 | -82.6028
4.6915 | -82.6024
4.6918 | -82.6016
4.6922 | -82.6012
4.6926 | -82.6009
4.6929 | -82.6013
4.6933 | -82.5999
4.6936 | -82.5994
4.694 | -82.5994
4.6944 | -82.6009
4.6947 | -82.6022
4.6951 | -82.602
4.6955 | -82.6014
4.6958 | -82.6017
4.6962 | -82.6028
4.6965 | -82.6018
4.6969 | -82.6002
4.6973 | -82.5975
4.6976 | -82.5967
4.698 | -82.596
4.6983 | -82.5958
4.6987 | -82.5958
4.6991 | -82.5956
4.6994 | -82.5949
4.6998 | -82.5958
4.7002 | -82.5966
4.7005 | -82.5965
4.7009 | -82.5959
4.7012 | -82.5955
4.7016 | -82.5942
4.702 | -82.5942
4.7023 | -82.5939
4.7027 | -82.5935
4.7031 | -82.5929
4.7034 | -82.5926
4.7038 | -82.5919
4.7041 | -82.5915
4.7045 | -82.5914
4.7049 | -82.5911
4.7052 | -82.5909
4.7056 | -82.5905
4.7059 | -82.5893
4.7063 | -82.5884
4.7067 | -82.5879
4.707 | -82.5871
4.7074 | -82.5865
4.7078 | -82.5862
4.7081 | -82.5857
4.7085 | -82.5849
4.7088 | -82.5844
4.7092 | -82.584
4.7096 | -82.5838
4.7099 | -82.5838
4.7103 | -82.5822
4.7107 | -82.5822
4.711 | -82.5816
4.7114 | -82.5812
4.7117 | -82.5817
4.7121 | -82.5817
4.7125 | -82.5814
4.7128 | -82.5809
4.7132 | -82.5805
4.7135 | -82.58
4.7139 | -82.5797
4.7143 | -82.5794
4.7146 | -82.5791
4.715 | -82.5786
4.7154 | -82.5781
4.7157 | -82.5764
4.7161 | -82.5765
4.7164 | -82.5766
4.7168 | -82.5759
4.7172 | -82.5751
4.7175 | -82.5743
4.7179 | -82.5749
4.7182 | -82.5737
4.7186 | -82.5741
4.719 | -82.5737
4.7193 | -82.5722
4.7197 | -82.5695
4.7201 | -82.5679
4.7204 | -82.5677
4.7208 | -82.5688
4.7211 | -82.5678
4.7215 | -82.5671
4.7219 | -82.5667
4.7222 | -82.5668
4.7226 | -82.5679
4.723 | -82.5687
4.7233 | -82.5689
4.7237 | -82.5684
4.724 | -82.5677
4.7244 | -82.5669
4.7248 | -82.5663
4.7251 | -82.5659
4.7255 | -82.5673
4.7258 | -82.5694
4.7262 | -82.5696
4.7266 | -82.5699
4.7269 | -82.5695
4.7273 | -82.5692
4.7277 | -82.5682
4.728 | -82.567
4.7284 | -82.5658
4.7287 | -82.5661
4.7291 | -82.5658
4.7295 | -82.5654
4.7298 | -82.5649
4.7302 | -82.5648
4.7305 | -82.5646
4.7309 | -82.5647
4.7313 | -82.5643
4.7316 | -82.5638
4.732 | -82.5628
4.7324 | -82.5625
4.7327 | -82.5618
4.7331 | -82.5611
4.7334 | -82.5613
4.7338 | -82.5617
4.7342 | -82.5613
4.7345 | -82.561
4.7349 | -82.5611
4.7353 | -82.5604
4.7356 | -82.5601
4.736 | -82.5591
4.7363 | -82.5587
4.7367 | -82.5583
4.7371 | -82.5601
4.7374 | -82.5605
4.7378 | -82.5606
4.7381 | -82.5601
4.7385 | -82.5584
4.7389 | -82.5575
4.7392 | -82.5582
4.7396 | -82.5573
4.74 | -82.5566
4.7403 | -82.5561
4.7407 | -82.556
4.741 | -82.5556
4.7414 | -82.5557
4.7418 | -82.5551
4.7421 | -82.555
4.7425 | -82.555
4.7428 | -82.5547
4.7432 | -82.5544
4.7436 | -82.5539
4.7439 | -82.5537
4.7443 | -82.553
4.7447 | -82.5528
4.745 | -82.5527
4.7454 | -82.5524
4.7457 | -82.5532
4.7461 | -82.5544
4.7465 | -82.5546
4.7468 | -82.557
4.7472 | -82.5597
4.7476 | -82.5588
4.7479 | -82.5575
4.7483 | -82.5557
4.7486 | -82.5544
4.749 | -82.5533
4.7494 | -82.5541
4.7497 | -82.5541
4.7501 | -82.5539
4.7504 | -82.5542
4.7508 | -82.5543
4.7512 | -82.5542
4.7515 | -82.5544
4.7519 | -82.555
4.7523 | -82.5551
4.7526 | -82.5549
4.753 | -82.5551
4.7533 | -82.5554
4.7537 | -82.5559
4.7541 | -82.5555
4.7544 | -82.5564
4.7548 | -82.5576
4.7552 | -82.5576
4.7555 | -82.5583
4.7559 | -82.5579
4.7562 | -82.5586
4.7566 | -82.5607
4.757 | -82.5617
4.7573 | -82.5627
4.7577 | -82.5628
4.758 | -82.563
4.7584 | -82.5637
4.7588 | -82.5638
4.7591 | -82.5625
4.7595 | -82.5621
4.7599 | -82.5625
4.7602 | -82.5653
4.7606 | -82.5695
4.7609 | -82.5717
4.7613 | -82.572
4.7617 | -82.5718
4.762 | -82.5716
4.7624 | -82.5714
4.7627 | -82.5708
4.7631 | -82.5697
4.7635 | -82.5684
4.7638 | -82.5684
4.7642 | -82.5679
4.7646 | -82.5677
4.7649 | -82.5672
4.7653 | -82.5668
4.7656 | -82.5667
4.766 | -82.5674
4.7664 | -82.567
4.7667 | -82.5661
4.7671 | -82.566
4.7675 | -82.5665
4.7678 | -82.5662
4.7682 | -82.5663
4.7685 | -82.565
4.7689 | -82.5648
4.7693 | -82.5652
4.7696 | -82.5651
4.77 | -82.5654
4.7703 | -82.5673
4.7707 | -82.5675
4.7711 | -82.5678
4.7714 | -82.5682
4.7718 | -82.5685
4.7722 | -82.5693
4.7725 | -82.5692
4.7729 | -82.5694
4.7732 | -82.5686
4.7736 | -82.569
4.774 | -82.5691
4.7743 | -82.5695
4.7747 | -82.57
4.7751 | -82.5705
4.7754 | -82.571
4.7758 | -82.5717
4.7761 | -82.5715
4.7765 | -82.5728
4.7769 | -82.5754
4.7772 | -82.5767
4.7776 | -82.5772
4.7779 | -82.578
4.7783 | -82.5798
4.7787 | -82.5807
4.779 | -82.581
4.7794 | -82.5815
4.7798 | -82.5826
4.7801 | -82.5855
4.7805 | -82.5872
4.7808 | -82.5885
4.7812 | -82.5895
4.7816 | -82.5908
4.7819 | -82.591
4.7823 | -82.5913
4.7826 | -82.5913
4.783 | -82.5927
4.7834 | -82.5933
4.7837 | -82.593
4.7841 | -82.5934
4.7845 | -82.5943
4.7848 | -82.5951
4.7852 | -82.596
4.7855 | -82.5964
4.7859 | -82.5961
4.7863 | -82.5974
4.7866 | -82.5997
4.787 | -82.601
4.7874 | -82.6025
4.7877 | -82.6035
4.7881 | -82.6047
4.7884 | -82.6064
4.7888 | -82.6069
4.7892 | -82.6082
4.7895 | -82.6106
4.7899 | -82.6126
4.7902 | -82.6133
4.7906 | -82.6142
4.791 | -82.6156
4.7913 | -82.6178
4.7917 | -82.6189
4.7921 | -82.6198
4.7924 | -82.6205
4.7928 | -82.6217
4.7931 | -82.6219
4.7935 | -82.6222
4.7939 | -82.6226
4.7942 | -82.6245
4.7946 | -82.6269
4.7949 | -82.6289
4.7953 | -82.6302
4.7957 | -82.6317
4.796 | -82.6331
4.7964 | -82.634
4.7968 | -82.6346
4.7971 | -82.6351
4.7975 | -82.6347
4.7978 | -82.6344
4.7982 | -82.6354
4.7986 | -82.6363
4.7989 | -82.6368
4.7993 | -82.6378
4.7997 | -82.6388
4.8 | -82.6396
4.8004 | -82.641
4.8007 | -82.6412
4.8011 | -82.6424
4.8015 | -82.6447
4.8018 | -82.6478
4.8022 | -82.6491
4.8025 | -82.65
4.8029 | -82.6516
4.8033 | -82.6531
4.8036 | -82.6562
4.804 | -82.6604
4.8044 | -82.6647
4.8047 | -82.6668
4.8051 | -82.6682
4.8054 | -82.67
4.8058 | -82.6714
4.8062 | -82.6723
4.8065 | -82.6743
4.8069 | -82.6767
4.8072 | -82.6787
4.8076 | -82.681
4.808 | -82.684
4.8083 | -82.6861
4.8087 | -82.6876
4.8091 | -82.6895
4.8094 | -82.6918
4.8098 | -82.6939
4.8101 | -82.6959
4.8105 | -82.6968
4.8109 | -82.6984
4.8112 | -82.6992
4.8116 | -82.6997
4.812 | -82.6995
4.8123 | -82.7014
4.8127 | -82.7031
4.813 | -82.7051
4.8134 | -82.7064
4.8138 | -82.708
4.8141 | -82.7091
4.8145 | -82.7101
4.8148 | -82.7119
4.8152 | -82.7143
4.8156 | -82.7165
4.8159 | -82.7183
4.8163 | -82.72
4.8167 | -82.7229
4.817 | -82.7257
4.8174 | -82.7277
4.8177 | -82.73
4.8181 | -82.7318
4.8185 | -82.7325
4.8188 | -82.7334
4.8192 | -82.7353
4.8196 | -82.7376
4.8199 | -82.7391
4.8203 | -82.7399
4.8206 | -82.7411
4.821 | -82.7429
4.8214 | -82.7447
4.8217 | -82.7462
4.8221 | -82.7479
4.8224 | -82.7507
4.8228 | -82.7524
4.8232 | -82.7545
4.8235 | -82.7559
4.8239 | -82.7569
4.8243 | -82.7589
4.8246 | -82.7606
4.825 | -82.7623
4.8253 | -82.7643
4.8257 | -82.7658
4.8261 | -82.7676
4.8264 | -82.7696
4.8268 | -82.7711
4.8271 | -82.7723
4.8275 | -82.774
4.8279 | -82.7755
4.8282 | -82.7771
4.8286 | -82.7797
4.829 | -82.7813
4.8293 | -82.7834
4.8297 | -82.7853
4.83 | -82.7868
4.8304 | -82.7884
4.8308 | -82.7901
4.8311 | -82.7916
4.8315 | -82.7932
4.8319 | -82.7945
4.8322 | -82.7966
4.8326 | -82.8003
4.8329 | -82.8034
4.8333 | -82.805
4.8337 | -82.8064
4.834 | -82.8083
4.8344 | -82.8113
4.8347 | -82.8136
4.8351 | -82.8154
4.8355 | -82.8173
4.8358 | -82.8188
4.8362 | -82.8215
4.8366 | -82.8238
4.8369 | -82.8269
4.8373 | -82.8292
4.8376 | -82.8307
4.838 | -82.8324
4.8384 | -82.8354
4.8387 | -82.8382
4.8391 | -82.8405
4.8395 | -82.8425
4.8398 | -82.8448
4.8402 | -82.8468
4.8405 | -82.8476
4.8409 | -82.8493
4.8413 | -82.8514
4.8416 | -82.8536
4.842 | -82.8569
4.8423 | -82.8609
4.8427 | -82.8629
4.8431 | -82.8663
4.8434 | -82.8694
4.8438 | -82.8723
4.8442 | -82.8745
4.8445 | -82.8768
4.8449 | -82.8792
4.8452 | -82.8806
4.8456 | -82.8817
4.846 | -82.884
4.8463 | -82.8876
4.8467 | -82.8904
4.847 | -82.8925
4.8474 | -82.895
4.8478 | -82.8973
4.8481 | -82.8996
4.8485 | -82.9017
4.8489 | -82.9037
4.8492 | -82.9059
4.8496 | -82.9083
4.8499 | -82.9112
4.8503 | -82.9138
4.8507 | -82.9163
4.851 | -82.9189
4.8514 | -82.9206
4.8518 | -82.9216
4.8521 | -82.9232
4.8525 | -82.9246
4.8528 | -82.9267
4.8532 | -82.9307
4.8536 | -82.9331
4.8539 | -82.9316
4.8543 | -82.9331
4.8546 | -82.9357
4.855 | -82.9383
4.8554 | -82.9403
4.8557 | -82.9421
4.8561 | -82.9443
4.8565 | -82.9465
4.8568 | -82.9484
4.8572 | -82.9509
4.8575 | -82.9538
4.8579 | -82.9564
4.8583 | -82.9582
4.8586 | -82.9594
4.859 | -82.9611
4.8594 | -82.9634
4.8597 | -82.9658
4.8601 | -82.9678
4.8604 | -82.9692
4.8608 | -82.9715
4.8612 | -82.9735
4.8615 | -82.9766
4.8619 | -82.9804
4.8622 | -82.9815
4.8626 | -82.9833
4.863 | -82.9854
4.8633 | -82.9872
4.8637 | -82.9898
4.8641 | -82.9918
4.8644 | -82.9936
4.8648 | -82.9959
4.8651 | -82.9997
4.8655 | -83.003
4.8659 | -83.0046
4.8662 | -83.0059
4.8666 | -83.008
4.8669 | -83.0103
4.8673 | -83.0123
4.8677 | -83.0147
4.868 | -83.0163
4.8684 | -83.0169
4.8688 | -83.0176
4.8691 | -83.0189
4.8695 | -83.0209
4.8698 | -83.023
4.8702 | -83.0254
4.8706 | -83.0287
4.8709 | -83.0307
4.8713 | -83.0323
4.8717 | -83.0342
4.872 | -83.0377
4.8724 | -83.0408
4.8727 | -83.0435
4.8731 | -83.0457
4.8735 | -83.0491
4.8738 | -83.0521
4.8742 | -83.0542
4.8745 | -83.0562
4.8749 | -83.0588
4.8753 | -83.0603
4.8756 | -83.0626
4.876 | -83.0648
4.8764 | -83.0665
4.8767 | -83.0689
4.8771 | -83.0719
4.8774 | -83.0743
4.8778 | -83.0758
4.8782 | -83.0775
4.8785 | -83.079
4.8789 | -83.081
4.8792 | -83.0826
4.8796 | -83.0846
4.88 | -83.0869
4.8803 | -83.0891
4.8807 | -83.0915
4.8811 | -83.0935
4.8814 | -83.095
4.8818 | -83.0971
4.8821 | -83.0989
4.8825 | -83.1006
4.8829 | -83.1026
4.8832 | -83.1046
4.8836 | -83.1067
4.884 | -83.1076
4.8843 | -83.1083
4.8847 | -83.1082
4.885 | -83.1081
4.8854 | -83.1087
4.8858 | -83.109
4.8861 | -83.1112
4.8865 | -83.1151
4.8868 | -83.1172
4.8872 | -83.1177
4.8876 | -83.1188
4.8879 | -83.1201
4.8883 | -83.1223
4.8887 | -83.1236
4.889 | -83.1225
4.8894 | -83.1239
4.8897 | -83.1263
4.8901 | -83.1277
4.8905 | -83.1283
4.8908 | -83.1293
4.8912 | -83.1307
4.8915 | -83.1327
4.8919 | -83.1341
4.8923 | -83.136
4.8926 | -83.1379
4.893 | -83.1385
4.8934 | -83.14
4.8937 | -83.1427
4.8941 | -83.1445
4.8944 | -83.1461
4.8948 | -83.1488
4.8952 | -83.1493
4.8955 | -83.1503
4.8959 | -83.1516
4.8963 | -83.1544
4.8966 | -83.1563
4.897 | -83.1576
4.8973 | -83.1588
4.8977 | -83.1608
4.8981 | -83.1629
4.8984 | -83.1644
4.8988 | -83.1658
4.8991 | -83.1679
4.8995 | -83.1694
4.8999 | -83.1714
4.9002 | -83.1735
4.9006 | -83.1747
4.901 | -83.1764
4.9013 | -83.1786
4.9017 | -83.1803
4.902 | -83.1823
4.9024 | -83.184
4.9028 | -83.1857
4.9031 | -83.1869
4.9035 | -83.1881
4.9039 | -83.1891
4.9042 | -83.1905
4.9046 | -83.1921
4.9049 | -83.1926
4.9053 | -83.1936
4.9057 | -83.1953
4.906 | -83.1961
4.9064 | -83.1973
4.9067 | -83.1986
4.9071 | -83.1999
4.9075 | -83.2011
4.9078 | -83.2022
4.9082 | -83.2036
4.9086 | -83.2042
4.9089 | -83.204
4.9093 | -83.205
4.9096 | -83.2074
4.91 | -83.2093
4.9104 | -83.2116
4.9107 | -83.2136
4.9111 | -83.2153
4.9114 | -83.2168
4.9118 | -83.2174
4.9122 | -83.2179
4.9125 | -83.2195
4.9129 | -83.222
4.9133 | -83.2237
4.9136 | -83.2254
4.914 | -83.2269
4.9143 | -83.2282
4.9147 | -83.2294
4.9151 | -83.2306
4.9154 | -83.2319
4.9158 | -83.2332
4.9162 | -83.2346
4.9165 | -83.2354
4.9169 | -83.2367
4.9172 | -83.2403
4.9176 | -83.2451
4.918 | -83.2487
4.9183 | -83.2502
4.9187 | -83.2525
4.919 | -83.2564
4.9194 | -83.2592
4.9198 | -83.2603
4.9201 | -83.2627
4.9205 | -83.2641
4.9209 | -83.2652
4.9212 | -83.267
4.9216 | -83.2695
4.9219 | -83.2711
4.9223 | -83.2724
4.9227 | -83.2739
4.923 | -83.2759
4.9234 | -83.2771
4.9238 | -83.2778
4.9241 | -83.2792
4.9245 | -83.2814
4.9248 | -83.2835
4.9252 | -83.2855
4.9256 | -83.2859
4.9259 | -83.2865
4.9263 | -83.2877
4.9266 | -83.2905
4.927 | -83.2918
4.9274 | -83.2936
4.9277 | -83.2948
4.9281 | -83.2973
4.9285 | -83.2986
4.9288 | -83.2993
4.9292 | -83.3002
4.9295 | -83.3009
4.9299 | -83.3028
4.9303 | -83.304
4.9306 | -83.3046
4.931 | -83.305
4.9313 | -83.3065
4.9317 | -83.3089
4.9321 | -83.3099
4.9324 | -83.3107
4.9328 | -83.3116
4.9332 | -83.3133
4.9335 | -83.3158
4.9339 | -83.3192
4.9342 | -83.3239
4.9346 | -83.3264
4.935 | -83.3289
4.9353 | -83.33
4.9357 | -83.3309
4.9361 | -83.3321
4.9364 | -83.3336
4.9368 | -83.336
4.9371 | -83.3357
4.9375 | -83.3355
4.9379 | -83.3373
4.9382 | -83.3393
4.9386 | -83.3396
4.9389 | -83.3409
4.9393 | -83.3417
4.9397 | -83.3418
4.94 | -83.3424
4.9404 | -83.3444
4.9408 | -83.3459
4.9411 | -83.3464
4.9415 | -83.3461
4.9418 | -83.3464
4.9422 | -83.3462
4.9426 | -83.3467
4.9429 | -83.3473
4.9433 | -83.3474
4.9436 | -83.349
4.944 | -83.3501
4.9444 | -83.3504
4.9447 | -83.3508
4.9451 | -83.3517
4.9455 | -83.3524
4.9458 | -83.3533
4.9462 | -83.3549
4.9465 | -83.3545
4.9469 | -83.3541
4.9473 | -83.3547
4.9476 | -83.356
4.948 | -83.3573
4.9484 | -83.3582
4.9487 | -83.3588
4.9491 | -83.3595
4.9494 | -83.36
4.9498 | -83.3607
4.9502 | -83.3613
4.9505 | -83.362
4.9509 | -83.3632
4.9512 | -83.365
4.9516 | -83.3669
4.952 | -83.3678
4.9523 | -83.3689
4.9527 | -83.3701
4.9531 | -83.3705
4.9534 | -83.3707
4.9538 | -83.3712
4.9541 | -83.3717
4.9545 | -83.3722
4.9549 | -83.3716
4.9552 | -83.3719
4.9556 | -83.3726
4.956 | -83.3732
4.9563 | -83.374
4.9567 | -83.3749
4.957 | -83.3755
4.9574 | -83.3762
4.9578 | -83.3768
4.9581 | -83.3781
4.9585 | -83.3801
4.9588 | -83.383
4.9592 | -83.3842
4.9596 | -83.3848
4.9599 | -83.3852
4.9603 | -83.3852
4.9607 | -83.3849
4.961 | -83.3858
4.9614 | -83.3861
4.9617 | -83.3868
4.9621 | -83.3875
4.9625 | -83.3885
4.9628 | -83.3895
4.9632 | -83.3902
4.9635 | -83.3903
4.9639 | -83.3913
4.9643 | -83.3918
4.9646 | -83.3909
4.965 | -83.3908
4.9654 | -83.3914
4.9657 | -83.3914
4.9661 | -83.3914
4.9664 | -83.3915
4.9668 | -83.3929
4.9672 | -83.394
4.9675 | -83.3946
4.9679 | -83.3953
4.9683 | -83.3958
4.9686 | -83.3957
4.969 | -83.3953
4.9693 | -83.396
4.9697 | -83.3968
4.9701 | -83.3972
4.9704 | -83.3981
4.9708 | -83.3995
4.9711 | -83.4005
4.9715 | -83.4022
4.9719 | -83.4049
4.9722 | -83.4066
4.9726 | -83.4073
4.973 | -83.4075
4.9733 | -83.4081
4.9737 | -83.4096
4.974 | -83.413
4.9744 | -83.4152
4.9748 | -83.4156
4.9751 | -83.4163
4.9755 | -83.417
4.9759 | -83.4185
4.9762 | -83.4196
4.9766 | -83.4201
4.9769 | -83.4206
4.9773 | -83.4191
4.9777 | -83.4186
4.978 | -83.4202
4.9784 | -83.422
4.9787 | -83.4236
4.9791 | -83.4246
4.9795 | -83.4272
4.9798 | -83.428
4.9802 | -83.4272
4.9806 | -83.4272
4.9809 | -83.4283
4.9813 | -83.4293
4.9816 | -83.4295
4.982 | -83.4302
4.9824 | -83.4311
4.9827 | -83.432
4.9831 | -83.432
4.9834 | -83.4329
4.9838 | -83.4337
4.9842 | -83.4338
4.9845 | -83.4337
4.9849 | -83.434
4.9853 | -83.435
4.9856 | -83.4353
4.986 | -83.4361
4.9863 | -83.4369
4.9867 | -83.4372
4.9871 | -83.4372
4.9874 | -83.4376
4.9878 | -83.4381
4.9882 | -83.438
4.9885 | -83.4387
4.9889 | -83.4396
4.9892 | -83.4394
4.9896 | -83.4405
4.99 | -83.4411
4.9903 | -83.442
4.9907 | -83.4446
4.991 | -83.4456
4.9914 | -83.4461
4.9918 | -83.447
4.9921 | -83.4473
4.9925 | -83.4481
4.9929 | -83.4506
4.9932 | -83.453
4.9936 | -83.4545
4.9939 | -83.4555
4.9943 | -83.456
4.9947 | -83.4568
4.995 | -83.4566
4.9954 | -83.4576
4.9957 | -83.4597
4.9961 | -83.4602
4.9965 | -83.461
4.9968 | -83.4614
4.9972 | -83.4628
4.9976 | -83.4648
4.9979 | -83.4658
4.9983 | -83.4667
4.9986 | -83.4666
4.999 | -83.468
4.9994 | -83.4689
4.9997 | -83.47
5.0001 | -83.4709
5.0005 | -83.4718
5.0008 | -83.4725
5.0012 | -83.4734
5.0015 | -83.4743
5.0019 | -83.475
5.0023 | -83.4778
5.0026 | -83.479
5.003 | -83.4802
5.0033 | -83.4805
5.0037 | -83.4811
5.0041 | -83.4814
5.0044 | -83.4803
5.0048 | -83.4806
5.0052 | -83.4808
5.0055 | -83.4808
5.0059 | -83.4813
5.0062 | -83.4828
5.0066 | -83.4851
5.007 | -83.486
5.0073 | -83.4871
5.0077 | -83.4883
5.008 | -83.4887
5.0084 | -83.489
5.0088 | -83.4896
5.0091 | -83.49
5.0095 | -83.4911
5.0099 | -83.4915
5.0102 | -83.4921
5.0106 | -83.4937
5.0109 | -83.4949
5.0113 | -83.4958
5.0117 | -83.4958
5.012 | -83.4958
5.0124 | -83.496
5.0128 | -83.4964
5.0131 | -83.4977
5.0135 | -83.4987
5.0138 | -83.501
5.0142 | -83.5033
5.0146 | -83.5044
5.0149 | -83.5062
5.0153 | -83.5072
5.0156 | -83.5075
5.016 | -83.5084
5.0164 | -83.5091
5.0167 | -83.5098
5.0171 | -83.5101
5.0175 | -83.5099
5.0178 | -83.5107
5.0182 | -83.5113
5.0185 | -83.5119
5.0189 | -83.5126
5.0193 | -83.5138
5.0196 | -83.5139
5.02 | -83.5147
5.0204 | -83.5157
5.0207 | -83.5166
5.0211 | -83.5185
5.0214 | -83.5191
5.0218 | -83.52
5.0222 | -83.5206
5.0225 | -83.5217
5.0229 | -83.5226
5.0232 | -83.5234
5.0236 | -83.5234
5.024 | -83.5239
5.0243 | -83.5258
5.0247 | -83.528
5.0251 | -83.5293
5.0254 | -83.5288
5.0258 | -83.5291
5.0261 | -83.5306
5.0265 | -83.5322
5.0269 | -83.5333
5.0272 | -83.5342
5.0276 | -83.5359
5.0279 | -83.5369
5.0283 | -83.5382
5.0287 | -83.5395
5.029 | -83.5405
5.0294 | -83.5419
5.0298 | -83.5428
5.0301 | -83.5437
5.0305 | -83.5444
5.0308 | -83.5467
5.0312 | -83.5486
5.0316 | -83.55
5.0319 | -83.5509
5.0323 | -83.5521
5.0327 | -83.5536
5.033 | -83.5562
5.0334 | -83.5574
5.0337 | -83.5585
5.0341 | -83.5596
5.0345 | -83.5612
5.0348 | -83.5636
5.0352 | -83.5662
5.0355 | -83.5683
5.0359 | -83.5703
5.0363 | -83.5711
5.0366 | -83.5722
5.037 | -83.5732
5.0374 | -83.5738
5.0377 | -83.5758
5.0381 | -83.577
5.0384 | -83.5781
5.0388 | -83.5792
5.0392 | -83.5802
5.0395 | -83.5812
5.0399 | -83.5839
5.0403 | -83.5871
5.0406 | -83.5882
5.041 | -83.5886
5.0413 | -83.5905
5.0417 | -83.5913
5.0421 | -83.5919
5.0424 | -83.5927
5.0428 | -83.5934
5.0431 | -83.5944
5.0435 | -83.5954
5.0439 | -83.5961
5.0442 | -83.5973
5.0446 | -83.5995
5.045 | -83.6018
5.0453 | -83.6031
5.0457 | -83.6041
5.046 | -83.6053
5.0464 | -83.607
5.0468 | -83.6084
5.0471 | -83.6097
5.0475 | -83.6114
5.0478 | -83.6129
5.0482 | -83.6143
5.0486 | -83.6144
5.0489 | -83.6141
5.0493 | -83.6168
5.0497 | -83.6188
5.05 | -83.6204
5.0504 | -83.6226
5.0507 | -83.6241
5.0511 | -83.6256
5.0515 | -83.6273
5.0518 | -83.6287
5.0522 | -83.6313
5.0526 | -83.6327
5.0529 | -83.634
5.0533 | -83.6355
5.0536 | -83.6366
5.054 | -83.6379
5.0544 | -83.639
5.0547 | -83.6401
5.0551 | -83.6431
5.0554 | -83.6449
5.0558 | -83.6467
5.0562 | -83.6488
5.0565 | -83.6507
5.0569 | -83.6529
5.0573 | -83.6549
5.0576 | -83.6573
5.058 | -83.6597
5.0583 | -83.6616
5.0587 | -83.6633
5.0591 | -83.6659
5.0594 | -83.67
5.0598 | -83.6728
5.0601 | -83.6742
5.0605 | -83.6752
5.0609 | -83.677
5.0612 | -83.6793
5.0616 | -83.6814
5.062 | -83.6835
5.0623 | -83.685
5.0627 | -83.6875
5.063 | -83.6887
5.0634 | -83.6904
5.0638 | -83.6913
5.0641 | -83.6933
5.0645 | -83.6951
5.0649 | -83.6969
5.0652 | -83.6977
5.0656 | -83.6988
5.0659 | -83.7005
5.0663 | -83.7028
5.0667 | -83.7043
5.067 | -83.7037
5.0674 | -83.7033
5.0677 | -83.7055
5.0681 | -83.7079
5.0685 | -83.7102
5.0688 | -83.7117
5.0692 | -83.7137
5.0696 | -83.7149
5.0699 | -83.7146
5.0703 | -83.7158
5.0706 | -83.7187
5.071 | -83.7198
5.0714 | -83.7201
5.0717 | -83.7208
5.0721 | -83.7217
5.0725 | -83.7234
5.0728 | -83.7264
5.0732 | -83.7283
5.0735 | -83.7302
5.0739 | -83.732
5.0743 | -83.7342
5.0746 | -83.7358
5.075 | -83.7374
5.0753 | -83.7383
5.0757 | -83.7394
5.0761 | -83.7412
5.0764 | -83.7426
5.0768 | -83.7453
5.0772 | -83.7471
5.0775 | -83.7495
5.0779 | -83.7524
5.0782 | -83.7554
5.0786 | -83.7562
5.079 | -83.7572
5.0793 | -83.7596
5.0797 | -83.7599
5.08 | -83.761
5.0804 | -83.7628
5.0808 | -83.7648
5.0811 | -83.7668
5.0815 | -83.7687
5.0819 | -83.7708
5.0822 | -83.7723
5.0826 | -83.7732
5.0829 | -83.7749
5.0833 | -83.7766
5.0837 | -83.7791
5.084 | -83.7809
5.0844 | -83.782
5.0848 | -83.7826
5.0851 | -83.7847
5.0855 | -83.7865
5.0858 | -83.7882
5.0862 | -83.7897
5.0866 | -83.7913
5.0869 | -83.7933
5.0873 | -83.7953
5.0876 | -83.7974
5.088 | -83.7993
5.0884 | -83.801
5.0887 | -83.8031
5.0891 | -83.8053
5.0895 | -83.8077
5.0898 | -83.8114
5.0902 | -83.8147
5.0905 | -83.8165
5.0909 | -83.818
5.0913 | -83.8197
5.0916 | -83.8211
5.092 | -83.8211
5.0923 | -83.8206
5.0927 | -83.821
5.0931 | -83.8217
5.0934 | -83.8227
5.0938 | -83.825
5.0942 | -83.8265
5.0945 | -83.8282
5.0949 | -83.8293
5.0952 | -83.8293
5.0956 | -83.8303
5.096 | -83.8321
5.0963 | -83.8345
5.0967 | -83.8364
5.0971 | -83.8386
5.0974 | -83.8413
5.0978 | -83.8436
5.0981 | -83.8455
5.0985 | -83.847
5.0989 | -83.8486
5.0992 | -83.8506
5.0996 | -83.8529
5.0999 | -83.8545
5.1003 | -83.8554
5.1007 | -83.8558
5.101 | -83.8571
5.1014 | -83.8602
5.1018 | -83.863
5.1021 | -83.8653
5.1025 | -83.8679
5.1028 | -83.8706
5.1032 | -83.8726
5.1036 | -83.8742
5.1039 | -83.8759
5.1043 | -83.8775
5.1047 | -83.8784
5.105 | -83.8801
5.1054 | -83.882
5.1057 | -83.8843
5.1061 | -83.8867
5.1065 | -83.8894
5.1068 | -83.8914
5.1072 | -83.8936
5.1075 | -83.8945
5.1079 | -83.8962
5.1083 | -83.8982
5.1086 | -83.8993
5.109 | -83.9002
5.1094 | -83.9023
5.1097 | -83.9042
5.1101 | -83.9059
5.1104 | -83.9077
5.1108 | -83.9101
5.1112 | -83.9118
5.1115 | -83.9152
5.1119 | -83.9172
5.1122 | -83.9189
5.1126 | -83.9206
5.113 | -83.9221
5.1133 | -83.9237
5.1137 | -83.9254
5.1141 | -83.927
5.1144 | -83.9286
5.1148 | -83.9299
5.1151 | -83.9315
5.1155 | -83.9331
5.1159 | -83.9348
5.1162 | -83.9367
5.1166 | -83.9394
5.117 | -83.9427
5.1173 | -83.9444
5.1177 | -83.9458
5.118 | -83.9461
5.1184 | -83.9471
5.1188 | -83.9488
5.1191 | -83.9506
5.1195 | -83.9525
5.1198 | -83.9547
5.1202 | -83.9565
5.1206 | -83.9582
5.1209 | -83.9599
5.1213 | -83.9613
5.1217 | -83.9622
5.122 | -83.9619
5.1224 | -83.9631
5.1227 | -83.9654
5.1231 | -83.9667
5.1235 | -83.967
5.1238 | -83.9684
5.1242 | -83.9698
5.1246 | -83.9715
5.1249 | -83.9738
5.1253 | -83.9742
5.1256 | -83.9749
5.126 | -83.9757
5.1264 | -83.9772
5.1267 | -83.9793
5.1271 | -83.9812
5.1274 | -83.9817
5.1278 | -83.9837
5.1282 | -83.9855
5.1285 | -83.9861
5.1289 | -83.9876
5.1293 | -83.989
5.1296 | -83.9901
5.13 | -83.9913
5.1303 | -83.994
5.1307 | -83.9956
5.1311 | -83.9973
5.1314 | -83.9988
5.1318 | -84.0004
5.1321 | -84.0025
5.1325 | -84.0036
5.1329 | -84.0053
5.1332 | -84.009
5.1336 | -84.011
5.134 | -84.0125
5.1343 | -84.0136
5.1347 | -84.0152
5.135 | -84.0165
5.1354 | -84.0174
5.1358 | -84.0188
5.1361 | -84.0206
5.1365 | -84.0223
5.1369 | -84.024
5.1372 | -84.0256
5.1376 | -84.0273
5.1379 | -84.0286
5.1383 | -84.0307
5.1387 | -84.0319
5.139 | -84.0339
5.1394 | -84.0355
5.1397 | -84.0368
5.1401 | -84.0374
5.1405 | -84.0385
5.1408 | -84.0399
5.1412 | -84.0412
5.1416 | -84.0426
5.1419 | -84.0443
5.1423 | -84.0461
5.1426 | -84.0494
5.143 | -84.0525
5.1434 | -84.0531
5.1437 | -84.0536
5.1441 | -84.0555
5.1444 | -84.0568
5.1448 | -84.058
5.1452 | -84.0599
5.1455 | -84.0613
5.1459 | -84.0643
5.1463 | -84.0658
5.1466 | -84.0677
5.147 | -84.0689
5.1473 | -84.0693
5.1477 | -84.0701
5.1481 | -84.0718
5.1484 | -84.0737
5.1488 | -84.0747
5.1492 | -84.0762
5.1495 | -84.0773
5.1499 | -84.0786
5.1502 | -84.0796
5.1506 | -84.0809
5.151 | -84.0826
5.1513 | -84.084
5.1517 | -84.0857
5.152 | -84.0874
5.1524 | -84.0884
5.1528 | -84.0893
5.1531 | -84.0902
5.1535 | -84.0917
5.1539 | -84.0928
5.1542 | -84.0949
5.1546 | -84.0962
5.1549 | -84.0974
5.1553 | -84.0978
5.1557 | -84.0984
5.156 | -84.0998
5.1564 | -84.1019
5.1567 | -84.1026
5.1571 | -84.1036
5.1575 | -84.106
5.1578 | -84.1091
5.1582 | -84.1127
5.1586 | -84.116
5.1589 | -84.1186
5.1593 | -84.1203
5.1596 | -84.1213
5.16 | -84.1226
5.1604 | -84.1238
5.1607 | -84.1256
5.1611 | -84.1272
5.1615 | -84.1283
5.1618 | -84.1281
5.1622 | -84.1283
5.1625 | -84.1292
5.1629 | -84.1307
5.1633 | -84.1315
5.1636 | -84.1329
5.164 | -84.1337
5.1643 | -84.1335
5.1647 | -84.1344
5.1651 | -84.1369
5.1654 | -84.1388
5.1658 | -84.1398
5.1662 | -84.1405
5.1665 | -84.1415
5.1669 | -84.1424
5.1672 | -84.145
5.1676 | -84.1476
5.168 | -84.1484
5.1683 | -84.1492
5.1687 | -84.1495
5.1691 | -84.151
5.1694 | -84.1524
5.1698 | -84.1537
5.1701 | -84.1549
5.1705 | -84.1559
5.1709 | -84.1574
5.1712 | -84.1591
5.1716 | -84.16
5.1719 | -84.1613
5.1723 | -84.1631
5.1727 | -84.164
5.173 | -84.1663
5.1734 | -84.1667
5.1738 | -84.1672
5.1741 | -84.1684
5.1745 | -84.1706
5.1748 | -84.1717
5.1752 | -84.1727
5.1756 | -84.174
5.1759 | -84.1752
5.1763 | -84.1757
5.1766 | -84.176
5.177 | -84.1761
5.1774 | -84.1774
5.1777 | -84.1788
5.1781 | -84.1801
5.1785 | -84.1809
5.1788 | -84.1815
5.1792 | -84.1821
5.1795 | -84.1838
5.1799 | -84.1855
5.1803 | -84.1861
5.1806 | -84.1868
5.181 | -84.1872
5.1814 | -84.1888
5.1817 | -84.1902
5.1821 | -84.1913
5.1824 | -84.1923
5.1828 | -84.193
5.1832 | -84.1938
5.1835 | -84.195
5.1839 | -84.196
5.1842 | -84.1963
5.1846 | -84.1972
5.185 | -84.1975
5.1853 | -84.1996
5.1857 | -84.2018
5.1861 | -84.2036
5.1864 | -84.2047
5.1868 | -84.2057
5.1871 | -84.2064
5.1875 | -84.2072
5.1879 | -84.2082
5.1882 | -84.2082
5.1886 | -84.2079
5.189 | -84.2087
5.1893 | -84.2085
5.1897 | -84.2088
5.19 | -84.2097
5.1904 | -84.2106
5.1908 | -84.2121
5.1911 | -84.2135
5.1915 | -84.2143
5.1918 | -84.2148
5.1922 | -84.214
5.1926 | -84.2112
5.1929 | -84.2109
5.1933 | -84.2114
5.1937 | -84.2119
5.194 | -84.2136
5.1944 | -84.2164
5.1947 | -84.2179
5.1951 | -84.2175
5.1955 | -84.218
5.1958 | -84.2184
5.1962 | -84.2189
5.1965 | -84.2193
5.1969 | -84.2197
5.1973 | -84.2201
5.1976 | -84.2201
5.198 | -84.2201
5.1984 | -84.2211
5.1987 | -84.2218
5.1991 | -84.2227
5.1994 | -84.2235
5.1998 | -84.2235
5.2002 | -84.2233
5.2005 | -84.2242
5.2009 | -84.2257
5.2013 | -84.2265
5.2016 | -84.2275
5.202 | -84.2289
5.2023 | -84.2297
5.2027 | -84.2306
5.2031 | -84.2332
5.2034 | -84.2361
5.2038 | -84.2368
5.2041 | -84.2364
5.2045 | -84.2359
5.2049 | -84.2363
5.2052 | -84.2371
5.2056 | -84.2377
5.206 | -84.2382
5.2063 | -84.2391
5.2067 | -84.2399
5.207 | -84.2403
5.2074 | -84.2407
5.2078 | -84.241
5.2081 | -84.2424
5.2085 | -84.2431
5.2088 | -84.2437
5.2092 | -84.2439
5.2096 | -84.2439
5.2099 | -84.244
5.2103 | -84.2452
5.2107 | -84.2452
5.211 | -84.2452
5.2114 | -84.2457
5.2117 | -84.2456
5.2121 | -84.2468
5.2125 | -84.2475
5.2128 | -84.247
5.2132 | -84.2467
5.2136 | -84.2473
5.2139 | -84.248
5.2143 | -84.249
5.2146 | -84.2499
5.215 | -84.25
5.2154 | -84.2487
5.2157 | -84.248
5.2161 | -84.2487
5.2164 | -84.2498
5.2168 | -84.2495
5.2172 | -84.2501
5.2175 | -84.2508
5.2179 | -84.251
5.2183 | -84.2512
5.2186 | -84.2535
5.219 | -84.2555
5.2193 | -84.2559
5.2197 | -84.257
5.2201 | -84.2574
5.2204 | -84.2568
5.2208 | -84.2568
5.2212 | -84.2572
5.2215 | -84.2575
5.2219 | -84.2574
5.2222 | -84.2576
5.2226 | -84.2577
5.223 | -84.2585
5.2233 | -84.2588
5.2237 | -84.2589
5.224 | -84.2578
5.2244 | -84.2571
5.2248 | -84.2573
5.2251 | -84.2576
5.2255 | -84.2579
5.2259 | -84.2583
5.2262 | -84.2593
5.2266 | -84.2604
5.2269 | -84.2611
5.2273 | -84.2605
5.2277 | -84.2579
5.228 | -84.2575
5.2284 | -84.2577
5.2287 | -84.2585
5.2291 | -84.2584
5.2295 | -84.2586
5.2298 | -84.2596
5.2302 | -84.259
5.2306 | -84.2589
5.2309 | -84.2592
5.2313 | -84.2591
5.2316 | -84.2571
5.232 | -84.2557
5.2324 | -84.2558
5.2327 | -84.2563
5.2331 | -84.2568
5.2335 | -84.2569
5.2338 | -84.2555
5.2342 | -84.2557
5.2345 | -84.2557
5.2349 | -84.2563
5.2353 | -84.258
5.2356 | -84.2596
5.236 | -84.2622
5.2363 | -84.264
5.2367 | -84.2653
5.2371 | -84.2674
5.2374 | -84.269
5.2378 | -84.2692
5.2382 | -84.2696
5.2385 | -84.2701
5.2389 | -84.2721
5.2392 | -84.274
5.2396 | -84.2749
5.24 | -84.2757
5.2403 | -84.2758
5.2407 | -84.2761
5.2411 | -84.2782
5.2414 | -84.2782
5.2418 | -84.2779
5.2421 | -84.278
5.2425 | -84.2774
5.2429 | -84.2759
5.2432 | -84.2762
5.2436 | -84.2759
5.2439 | -84.2755
5.2443 | -84.2749
5.2447 | -84.2751
5.245 | -84.2765
5.2454 | -84.277
5.2458 | -84.2769
5.2461 | -84.2774
5.2465 | -84.2779
5.2468 | -84.2782
5.2472 | -84.2791
5.2476 | -84.2782
5.2479 | -84.2787
5.2483 | -84.2791
5.2486 | -84.2784
5.249 | -84.2786
5.2494 | -84.2784
5.2497 | -84.2786
5.2501 | -84.2769
5.2505 | -84.2768
5.2508 | -84.2769
5.2512 | -84.2767
5.2515 | -84.2768
5.2519 | -84.2764
5.2523 | -84.2755
5.2526 | -84.2749
5.253 | -84.2748
5.2534 | -84.2752
5.2537 | -84.2749
5.2541 | -84.2737
5.2544 | -84.273
5.2548 | -84.2734
5.2552 | -84.2738
5.2555 | -84.2737
5.2559 | -84.2737
5.2562 | -84.274
5.2566 | -84.2746
5.257 | -84.275
5.2573 | -84.2751
5.2577 | -84.2767
5.2581 | -84.2773
5.2584 | -84.2772
5.2588 | -84.2775
5.2591 | -84.2778
5.2595 | -84.2779
5.2599 | -84.2781
5.2602 | -84.278
5.2606 | -84.2782
5.2609 | -84.2788
5.2613 | -84.2788
5.2617 | -84.2788
5.262 | -84.28
5.2624 | -84.2821
5.2628 | -84.2834
5.2631 | -84.2833
5.2635 | -84.283
5.2638 | -84.2831
5.2642 | -84.2836
5.2646 | -84.2836
5.2649 | -84.2836
5.2653 | -84.2837
5.2657 | -84.2844
5.266 | -84.2842
5.2664 | -84.2843
5.2667 | -84.284
5.2671 | -84.2839
5.2675 | -84.2831
5.2678 | -84.2821
5.2682 | -84.2818
5.2685 | -84.2818
5.2689 | -84.2822
5.2693 | -84.2825
5.2696 | -84.2811
5.27 | -84.2801
5.2704 | -84.2798
5.2707 | -84.2797
5.2711 | -84.2795
5.2714 | -84.2785
5.2718 | -84.2786
5.2722 | -84.2783
5.2725 | -84.2778
5.2729 | -84.2777
5.2732 | -84.2777
5.2736 | -84.2771
5.274 | -84.276
5.2743 | -84.2745
5.2747 | -84.2733
5.2751 | -84.2726
5.2754 | -84.272
5.2758 | -84.2709
5.2761 | -84.2692
5.2765 | -84.2682
5.2769 | -84.2685
5.2772 | -84.2683
5.2776 | -84.2678
5.278 | -84.2684
5.2783 | -84.2699
5.2787 | -84.2715
5.279 | -84.2726
5.2794 | -84.2721
5.2798 | -84.273
5.2801 | -84.2729
5.2805 | -84.272
5.2808 | -84.2716
5.2812 | -84.2715
5.2816 | -84.2714
5.2819 | -84.2721
5.2823 | -84.2721
5.2827 | -84.272
5.283 | -84.271
5.2834 | -84.2705
5.2837 | -84.2707
5.2841 | -84.2708
5.2845 | -84.2702
5.2848 | -84.2702
5.2852 | -84.2704
5.2856 | -84.2708
5.2859 | -84.2715
5.2863 | -84.2714
5.2866 | -84.2715
5.287 | -84.2713
5.2874 | -84.2713
5.2877 | -84.2714
5.2881 | -84.2709
5.2884 | -84.27
5.2888 | -84.27
5.2892 | -84.2704
5.2895 | -84.2704
5.2899 | -84.2697
5.2903 | -84.2695
5.2906 | -84.2695
5.291 | -84.2692
5.2913 | -84.2687
5.2917 | -84.269
5.2921 | -84.2707
5.2924 | -84.2714
5.2928 | -84.2727
5.2931 | -84.2717
5.2935 | -84.2704
5.2939 | -84.2702
5.2942 | -84.2711
5.2946 | -84.27
5.295 | -84.2677
5.2953 | -84.2672
5.2957 | -84.267
5.296 | -84.2671
5.2964 | -84.2671
5.2968 | -84.2668
5.2971 | -84.2664
5.2975 | -84.2679
5.2979 | -84.2694
5.2982 | -84.2699
5.2986 | -84.2705
5.2989 | -84.2712
5.2993 | -84.2707
5.2997 | -84.2699
5.3 | -84.2705
5.3004 | -84.2706
5.3007 | -84.2705
5.3011 | -84.2703
5.3015 | -84.27
5.3018 | -84.2685
5.3022 | -84.2692
5.3026 | -84.2689
5.3029 | -84.2684
5.3033 | -84.2681
5.3036 | -84.2677
5.304 | -84.2675
5.3044 | -84.2673
5.3047 | -84.2676
5.3051 | -84.2678
5.3055 | -84.2671
5.3058 | -84.2664
5.3062 | -84.266
5.3065 | -84.2654
5.3069 | -84.2651
5.3073 | -84.265
5.3076 | -84.2666
5.308 | -84.2674
5.3083 | -84.2674
5.3087 | -84.2665
5.3091 | -84.2659
5.3094 | -84.2668
5.3098 | -84.2674
5.3102 | -84.2674
5.3105 | -84.2671
5.3109 | -84.2669
5.3112 | -84.2681
5.3116 | -84.2686
5.312 | -84.2682
5.3123 | -84.2673
5.3127 | -84.266
5.313 | -84.2664
5.3134 | -84.2656
5.3138 | -84.2645
5.3141 | -84.2631
5.3145 | -84.2624
5.3149 | -84.2618
5.3152 | -84.2613
5.3156 | -84.2612
5.3159 | -84.2612
5.3163 | -84.2611
5.3167 | -84.2609
5.317 | -84.2617
5.3174 | -84.2629
5.3178 | -84.2635
5.3181 | -84.2629
5.3185 | -84.2632
5.3188 | -84.2627
5.3192 | -84.2621
5.3196 | -84.261
5.3199 | -84.259
5.3203 | -84.2582
5.3206 | -84.2585
5.321 | -84.2586
5.3214 | -84.2591
5.3217 | -84.2604
5.3221 | -84.261
5.3225 | -84.2607
5.3228 | -84.261
5.3232 | -84.262
5.3235 | -84.2605
5.3239 | -84.2602
5.3243 | -84.2591
5.3246 | -84.258
5.325 | -84.2575
5.3254 | -84.2574
5.3257 | -84.2575
5.3261 | -84.2578
5.3264 | -84.2586
5.3268 | -84.2593
5.3272 | -84.2589
5.3275 | -84.2588
5.3279 | -84.2586
5.3282 | -84.2582
5.3286 | -84.2578
5.329 | -84.2574
5.3293 | -84.2571
5.3297 | -84.2565
5.3301 | -84.2565
5.3304 | -84.257
5.3308 | -84.2576
5.3311 | -84.2569
5.3315 | -84.2566
5.3319 | -84.2563
5.3322 | -84.256
5.3326 | -84.256
5.3329 | -84.256
5.3333 | -84.2553
5.3337 | -84.2537
5.334 | -84.2534
5.3344 | -84.253
5.3348 | -84.2525
5.3351 | -84.2523
5.3355 | -84.2554
5.3358 | -84.2565
5.3362 | -84.2565
5.3366 | -84.2558
5.3369 | -84.2548
5.3373 | -84.2544
5.3377 | -84.2543
5.338 | -84.2534
5.3384 | -84.2532
5.3387 | -84.2537
5.3391 | -84.2534
5.3395 | -84.2529
5.3398 | -84.2522
5.3402 | -84.252
5.3405 | -84.2511
5.3409 | -84.2504
5.3413 | -84.2499
5.3416 | -84.25
5.342 | -84.2495
5.3424 | -84.2491
5.3427 | -84.2491
5.3431 | -84.2498
5.3434 | -84.2497
5.3438 | -84.2491
5.3442 | -84.2503
5.3445 | -84.2498
5.3449 | -84.249
5.3452 | -84.2488
5.3456 | -84.2489
5.346 | -84.2495
5.3463 | -84.2494
5.3467 | -84.2489
5.3471 | -84.2488
5.3474 | -84.2488
5.3478 | -84.2486
5.3481 | -84.2491
5.3485 | -84.2487
5.3489 | -84.2484
5.3492 | -84.2476
5.3496 | -84.2457
5.35 | -84.2445
5.3503 | -84.2446
5.3507 | -84.2441
5.351 | -84.2432
5.3514 | -84.2428
5.3518 | -84.2426
5.3521 | -84.2416
5.3525 | -84.2411
5.3528 | -84.2406
5.3532 | -84.24
5.3536 | -84.2395
5.3539 | -84.2388
5.3543 | -84.2394
5.3547 | -84.2386
5.355 | -84.2373
5.3554 | -84.2369
5.3557 | -84.2369
5.3561 | -84.2365
5.3565 | -84.2361
5.3568 | -84.2351
5.3572 | -84.2355
5.3575 | -84.2343
5.3579 | -84.2338
5.3583 | -84.2343
5.3586 | -84.2344
5.359 | -84.2335
5.3594 | -84.2327
5.3597 | -84.2326
5.3601 | -84.2321
5.3604 | -84.2317
5.3608 | -84.2312
5.3612 | -84.2308
5.3615 | -84.2299
5.3619 | -84.2296
5.3623 | -84.2293
5.3626 | -84.2277
5.363 | -84.2255
5.3633 | -84.224
5.3637 | -84.2235
5.3641 | -84.2235
5.3644 | -84.2246
5.3648 | -84.2247
5.3651 | -84.2252
5.3655 | -84.22
5.3659 | -84.2192
5.3662 | -84.2189
5.3666 | -84.2174
5.367 | -84.2168
5.3673 | -84.2164
5.3677 | -84.216
5.368 | -84.2152
5.3684 | -84.2146
5.3688 | -84.2144
5.3691 | -84.2141
5.3695 | -84.2137
5.3699 | -84.2133
5.3702 | -84.2131
5.3706 | -84.2123
5.3709 | -84.2115
5.3713 | -84.2113
5.3717 | -84.211
5.372 | -84.2099
5.3724 | -84.2101
5.3727 | -84.2108
5.3731 | -84.2104
5.3735 | -84.2097
5.3738 | -84.2093
5.3742 | -84.21
5.3746 | -84.2103
5.3749 | -84.2103
5.3753 | -84.2096
5.3756 | -84.209
5.376 | -84.2088
5.3764 | -84.2092
5.3767 | -84.21
5.3771 | -84.2086
5.3774 | -84.2077
5.3778 | -84.2069
5.3782 | -84.2061
5.3785 | -84.2066
5.3789 | -84.2058
5.3793 | -84.205
5.3796 | -84.2036
5.38 | -84.2017
5.3803 | -84.201
5.3807 | -84.2011
5.3811 | -84.2012
5.3814 | -84.2007
5.3818 | -84.1996
5.3822 | -84.1992
5.3825 | -84.1981
5.3829 | -84.1978
5.3832 | -84.1976
5.3836 | -84.197
5.384 | -84.1964
5.3843 | -84.1958
5.3847 | -84.1947
5.385 | -84.1935
5.3854 | -84.1919
5.3858 | -84.1907
5.3861 | -84.191
5.3865 | -84.1908
5.3869 | -84.1906
5.3872 | -84.1914
5.3876 | -84.1908
5.3879 | -84.1903
5.3883 | -84.1896
5.3887 | -84.189
5.389 | -84.1885
5.3894 | -84.1878
5.3898 | -84.1868
5.3901 | -84.1848
5.3905 | -84.1836
5.3908 | -84.183
5.3912 | -84.1827
5.3916 | -84.1821
5.3919 | -84.1819
5.3923 | -84.1818
5.3926 | -84.1813
5.393 | -84.1808
5.3934 | -84.1788
5.3937 | -84.1773
5.3941 | -84.1766
5.3945 | -84.1756
5.3948 | -84.1751
5.3952 | -84.1745
5.3955 | -84.1743
5.3959 | -84.174
5.3963 | -84.1734
5.3966 | -84.1729
5.397 | -84.1722
5.3973 | -84.172
5.3977 | -84.1719
5.3981 | -84.1721
5.3984 | -84.1717
5.3988 | -84.1717
5.3992 | -84.1716
5.3995 | -84.1705
5.3999 | -84.1693
5.4002 | -84.169
5.4006 | -84.1692
5.401 | -84.1686
5.4013 | -84.1683
5.4017 | -84.168
5.4021 | -84.1671
5.4024 | -84.1659
5.4028 | -84.1653
5.4031 | -84.1653
5.4035 | -84.1656
5.4039 | -84.1649
5.4042 | -84.165
5.4046 | -84.1656
5.4049 | -84.1651
5.4053 | -84.165
5.4057 | -84.1654
5.406 | -84.1649
5.4064 | -84.1641
5.4068 | -84.1635
5.4071 | -84.1636
5.4075 | -84.163
5.4078 | -84.162
5.4082 | -84.1611
5.4086 | -84.1608
5.4089 | -84.1608
5.4093 | -84.1597
5.4096 | -84.1589
5.41 | -84.1587
5.4104 | -84.1585
5.4107 | -84.1583
5.4111 | -84.1577
5.4115 | -84.1571
5.4118 | -84.158
5.4122 | -84.1592
5.4125 | -84.1583
5.4129 | -84.1573
5.4133 | -84.1563
5.4136 | -84.1561
5.414 | -84.1556
5.4144 | -84.1553
5.4147 | -84.1546
5.4151 | -84.1543
5.4154 | -84.1539
5.4158 | -84.1537
5.4162 | -84.1531
5.4165 | -84.1523
5.4169 | -84.1519
5.4172 | -84.152
5.4176 | -84.1516
5.418 | -84.1511
5.4183 | -84.151
5.4187 | -84.1512
5.4191 | -84.1511
5.4194 | -84.1509
5.4198 | -84.1507
5.4201 | -84.1505
5.4205 | -84.1505
5.4209 | -84.15
5.4212 | -84.1498
5.4216 | -84.1494
5.4219 | -84.1492
5.4223 | -84.1486
5.4227 | -84.1483
5.423 | -84.1483
5.4234 | -84.1481
5.4238 | -84.1483
5.4241 | -84.1482
5.4245 | -84.148
5.4248 | -84.148
5.4252 | -84.148
5.4256 | -84.148
5.4259 | -84.1477
5.4263 | -84.1478
5.4267 | -84.1474
5.427 | -84.1479
5.4274 | -84.1481
5.4277 | -84.1482
5.4281 | -84.1487
5.4285 | -84.1496
5.4288 | -84.1495
5.4292 | -84.1497
5.4295 | -84.1493
5.4299 | -84.1504
5.4303 | -84.1516
5.4306 | -84.1528
5.431 | -84.1527
5.4314 | -84.1532
5.4317 | -84.1564
5.4321 | -84.1586
5.4324 | -84.1589
5.4328 | -84.1584
5.4332 | -84.158
5.4335 | -84.1583
5.4339 | -84.1575
5.4343 | -84.1576
5.4346 | -84.159
5.435 | -84.1592
5.4353 | -84.1597
5.4357 | -84.1606
5.4361 | -84.1622
5.4364 | -84.1629
5.4368 | -84.1633
5.4371 | -84.1634
5.4375 | -84.1624
5.4379 | -84.1632
5.4382 | -84.1634
5.4386 | -84.1641
5.439 | -84.1644
5.4393 | -84.1651
5.4397 | -84.1654
5.44 | -84.1656
5.4404 | -84.1658
5.4408 | -84.1667
5.4411 | -84.167
5.4415 | -84.1677
5.4418 | -84.1681
5.4422 | -84.1677
5.4426 | -84.1677
5.4429 | -84.1679
5.4433 | -84.1683
5.4437 | -84.169
5.444 | -84.1711
5.4444 | -84.1723
5.4447 | -84.1731
5.4451 | -84.1754
5.4455 | -84.1762
5.4458 | -84.1757
5.4462 | -84.174
5.4466 | -84.1734
5.4469 | -84.1737
5.4473 | -84.1741
5.4476 | -84.1745
5.448 | -84.1756
5.4484 | -84.1768
5.4487 | -84.1779
5.4491 | -84.1784
5.4494 | -84.1788
5.4498 | -84.1794
5.4502 | -84.1796
5.4505 | -84.1805
5.4509 | -84.1817
5.4513 | -84.1819
5.4516 | -84.1817
5.452 | -84.1825
5.4523 | -84.1828
5.4527 | -84.1827
5.4531 | -84.1831
5.4534 | -84.1834
5.4538 | -84.1839
5.4542 | -84.1843
5.4545 | -84.1848
5.4549 | -84.1854
5.4552 | -84.1866
5.4556 | -84.1878
5.456 | -84.1882
5.4563 | -84.189
5.4567 | -84.1899
5.457 | -84.191
5.4574 | -84.1917
5.4578 | -84.1926
5.4581 | -84.1948
5.4585 | -84.1975
5.4589 | -84.1991
5.4592 | -84.2003
5.4596 | -84.202
5.4599 | -84.2028
5.4603 | -84.2034
5.4607 | -84.2041
5.461 | -84.2054
5.4614 | -84.2069
5.4617 | -84.2086
5.4621 | -84.2111
5.4625 | -84.2119
5.4628 | -84.2123
5.4632 | -84.2127
5.4636 | -84.2134
5.4639 | -84.2131
5.4643 | -84.2123
5.4646 | -84.2123
5.465 | -84.2134
5.4654 | -84.2144
5.4657 | -84.215
5.4661 | -84.2156
5.4665 | -84.216
5.4668 | -84.2175
5.4672 | -84.2192
5.4675 | -84.2214
5.4679 | -84.2232
5.4683 | -84.2242
5.4686 | -84.2254
5.469 | -84.2263
5.4693 | -84.2264
5.4697 | -84.2284
5.4701 | -84.2288
5.4704 | -84.2299
5.4708 | -84.2316
5.4712 | -84.2328
5.4715 | -84.2337
5.4719 | -84.2347
5.4722 | -84.2359
5.4726 | -84.2367
5.473 | -84.2364
5.4733 | -84.2363
5.4737 | -84.2375
5.4741 | -84.2386
5.4744 | -84.2391
5.4748 | -84.2399
5.4751 | -84.2415
5.4755 | -84.243
5.4759 | -84.2445
5.4762 | -84.246
5.4766 | -84.2471
5.4769 | -84.2477
5.4773 | -84.2485
5.4777 | -84.2496
5.478 | -84.2507
5.4784 | -84.2513
5.4788 | -84.2518
5.4791 | -84.2515
5.4795 | -84.2526
5.4798 | -84.2533
5.4802 | -84.2551
5.4806 | -84.2558
5.4809 | -84.257
5.4813 | -84.2575
5.4816 | -84.2577
5.482 | -84.2587
5.4824 | -84.2597
5.4827 | -84.2611
5.4831 | -84.2617
5.4835 | -84.2622
5.4838 | -84.2634
5.4842 | -84.2639
5.4845 | -84.2647
5.4849 | -84.2655
5.4853 | -84.2662
5.4856 | -84.2672
5.486 | -84.2688
5.4864 | -84.2712
5.4867 | -84.2727
5.4871 | -84.2737
5.4874 | -84.275
5.4878 | -84.2764
5.4882 | -84.2776
5.4885 | -84.2778
5.4889 | -84.2786
5.4892 | -84.28
5.4896 | -84.2813
5.49 | -84.2833
5.4903 | -84.2863
5.4907 | -84.2886
5.4911 | -84.2895
5.4914 | -84.2906
5.4918 | -84.2928
5.4921 | -84.2947
5.4925 | -84.2961
5.4929 | -84.2974
5.4932 | -84.2986
5.4936 | -84.3003
5.4939 | -84.3023
5.4943 | -84.3044
5.4947 | -84.3057
5.495 | -84.3079
5.4954 | -84.3099
5.4958 | -84.311
5.4961 | -84.3119
5.4965 | -84.3135
5.4968 | -84.3147
5.4972 | -84.3168
5.4976 | -84.3194
5.4979 | -84.3203
5.4983 | -84.3209
5.4987 | -84.3219
5.499 | -84.3244
5.4994 | -84.3262
5.4997 | -84.3276
5.5001 | -84.3289
5.5005 | -84.3297
5.5008 | -84.3306
5.5012 | -84.3328
5.5015 | -84.3334
5.5019 | -84.3343
5.5023 | -84.3354
5.5026 | -84.3362
5.503 | -84.3363
5.5034 | -84.3362
5.5037 | -84.3372
5.5041 | -84.3388
5.5044 | -84.3405
5.5048 | -84.342
5.5052 | -84.3433
5.5055 | -84.3448
5.5059 | -84.3459
5.5062 | -84.3469
5.5066 | -84.3475
5.507 | -84.3488
5.5073 | -84.3503
5.5077 | -84.3519
5.5081 | -84.3538
5.5084 | -84.3565
5.5088 | -84.3584
5.5091 | -84.3598
5.5095 | -84.3613
5.5099 | -84.3633
5.5102 | -84.3649
5.5106 | -84.3659
5.511 | -84.3667
5.5113 | -84.368
5.5117 | -84.3701
5.512 | -84.3719
5.5124 | -84.3739
5.5128 | -84.3751
5.5131 | -84.3754
5.5135 | -84.3772
5.5138 | -84.3793
5.5142 | -84.3802
5.5146 | -84.3815
5.5149 | -84.3824
5.5153 | -84.3839
5.5157 | -84.3855
5.516 | -84.3866
5.5164 | -84.3877
5.5167 | -84.3886
5.5171 | -84.3897
5.5175 | -84.3912
5.5178 | -84.3937
5.5182 | -84.3951
5.5186 | -84.3962
5.5189 | -84.3976
5.5193 | -84.3992
5.5196 | -84.4009
5.52 | -84.4024
5.5204 | -84.4041
5.5207 | -84.4062
5.5211 | -84.408
5.5214 | -84.4089
5.5218 | -84.4058
5.5222 | -84.406
5.5225 | -84.4076
5.5229 | -84.4095
5.5233 | -84.4112
5.5236 | -84.4133
5.524 | -84.4154
5.5243 | -84.4163
5.5247 | -84.4173
5.5251 | -84.4188
5.5254 | -84.4203
5.5258 | -84.4218
5.5261 | -84.4244
5.5265 | -84.4263
5.5269 | -84.4282
5.5272 | -84.4297
5.5276 | -84.4322
5.528 | -84.4334
5.5283 | -84.4353
5.5287 | -84.4363
5.529 | -84.4376
5.5294 | -84.4388
5.5298 | -84.44
5.5301 | -84.4414
5.5305 | -84.4431
5.5309 | -84.4446
5.5312 | -84.4457
5.5316 | -84.4466
5.5319 | -84.448
5.5323 | -84.4505
5.5327 | -84.4525
5.533 | -84.4544
5.5334 | -84.4573
5.5337 | -84.4596
5.5341 | -84.4604
5.5345 | -84.4615
5.5348 | -84.4625
5.5352 | -84.4638
5.5356 | -84.4653
5.5359 | -84.4681
5.5363 | -84.4707
5.5366 | -84.4733
5.537 | -84.4754
5.5374 | -84.4753
5.5377 | -84.4761
5.5381 | -84.4776
5.5385 | -84.4791
5.5388 | -84.4798
5.5392 | -84.4815
5.5395 | -84.4829
5.5399 | -84.4848
5.5403 | -84.4864
5.5406 | -84.4891
5.541 | -84.4904
5.5413 | -84.4917
5.5417 | -84.4933
5.5421 | -84.4948
5.5424 | -84.4961
5.5428 | -84.4978
5.5432 | -84.4997
5.5435 | -84.502
5.5439 | -84.5036
5.5442 | -84.5052
5.5446 | -84.5069
5.545 | -84.5086
5.5453 | -84.5102
5.5457 | -84.5131
5.546 | -84.5146
5.5464 | -84.5162
5.5468 | -84.5178
5.5471 | -84.5197
5.5475 | -84.5218
5.5479 | -84.523
5.5482 | -84.5234
5.5486 | -84.5255
5.5489 | -84.5278
5.5493 | -84.5305
5.5497 | -84.5319
5.55 | -84.5334
5.5504 | -84.5351
5.5508 | -84.5371
5.5511 | -84.5392
5.5515 | -84.542
5.5518 | -84.5444
5.5522 | -84.5464
5.5526 | -84.5484
5.5529 | -84.5504
5.5533 | -84.552
5.5536 | -84.5533
5.554 | -84.5553
5.5544 | -84.5574
5.5547 | -84.5592
5.5551 | -84.561
5.5555 | -84.5625
5.5558 | -84.5641
5.5562 | -84.5658
5.5565 | -84.5675
5.5569 | -84.5695
5.5573 | -84.5714
5.5576 | -84.5726
5.558 | -84.5731
5.5583 | -84.5745
5.5587 | -84.5754
5.5591 | -84.5768
5.5594 | -84.5784
5.5598 | -84.5799
5.5602 | -84.581
5.5605 | -84.5824
5.5609 | -84.5835
5.5612 | -84.585
5.5616 | -84.5874
5.562 | -84.5901
5.5623 | -84.5932
5.5627 | -84.5967
5.5631 | -84.5987
5.5634 | -84.6004
5.5638 | -84.6021
5.5641 | -84.6042
5.5645 | -84.6056
5.5649 | -84.6075
5.5652 | -84.6093
5.5656 | -84.6107
5.5659 | -84.6128
5.5663 | -84.6184
5.5667 | -84.621
5.567 | -84.6226
5.5674 | -84.624
5.5678 | -84.6259
5.5681 | -84.6269
5.5685 | -84.6286
5.5688 | -84.6312
5.5692 | -84.6327
5.5696 | -84.6342
5.5699 | -84.6353
5.5703 | -84.6368
5.5707 | -84.6378
5.571 | -84.6389
5.5714 | -84.641
5.5717 | -84.6416
5.5721 | -84.6406
5.5725 | -84.6406
5.5728 | -84.642
5.5732 | -84.6436
5.5735 | -84.6447
5.5739 | -84.6469
5.5743 | -84.648
5.5746 | -84.6488
5.575 | -84.6483
5.5754 | -84.6495
5.5757 | -84.6509
5.5761 | -84.6528
5.5764 | -84.6543
5.5768 | -84.6557
5.5772 | -84.6576
5.5775 | -84.659
5.5779 | -84.6605
5.5782 | -84.6626
5.5786 | -84.6651
5.579 | -84.6675
5.5793 | -84.6692
5.5797 | -84.6715
5.5801 | -84.6726
5.5804 | -84.6742
5.5808 | -84.6755
5.5811 | -84.6769
5.5815 | -84.6786
5.5819 | -84.6795
5.5822 | -84.6807
5.5826 | -84.6826
5.583 | -84.6837
5.5833 | -84.6852
5.5837 | -84.6866
5.584 | -84.6882
5.5844 | -84.69
5.5848 | -84.6902
5.5851 | -84.6914
5.5855 | -84.6925
5.5858 | -84.6941
5.5862 | -84.6953
5.5866 | -84.6955
5.5869 | -84.6972
5.5873 | -84.6994
5.5877 | -84.7008
5.588 | -84.7031
5.5884 | -84.7053
5.5887 | -84.7056
5.5891 | -84.706
5.5895 | -84.7075
5.5898 | -84.7088
5.5902 | -84.7101
5.5906 | -84.7123
5.5909 | -84.7151
5.5913 | -84.7168
5.5916 | -84.7182
5.592 | -84.7193
5.5924 | -84.7191
5.5927 | -84.7194
5.5931 | -84.7203
5.5934 | -84.7222
5.5938 | -84.7238
5.5942 | -84.7249
5.5945 | -84.7261
5.5949 | -84.7277
5.5953 | -84.7294
5.5956 | -84.7307
5.596 | -84.7314
5.5963 | -84.7333
5.5967 | -84.7345
5.5971 | -84.7352
5.5974 | -84.7363
5.5978 | -84.7383
5.5981 | -84.7382
5.5985 | -84.7378
5.5989 | -84.7386
5.5992 | -84.7397
5.5996 | -84.7411
5.6 | -84.7434
5.6003 | -84.7455
5.6007 | -84.7468
5.601 | -84.748
5.6014 | -84.7492
5.6018 | -84.7496
5.6021 | -84.7504
5.6025 | -84.7512
5.6029 | -84.7524
5.6032 | -84.7543
5.6036 | -84.7582
5.6039 | -84.7601
5.6043 | -84.7602
5.6047 | -84.7612
5.605 | -84.7627
5.6054 | -84.7622
5.6057 | -84.764
5.6061 | -84.7649
5.6065 | -84.7654
5.6068 | -84.7671
5.6072 | -84.769
5.6076 | -84.7698
5.6079 | -84.7702
5.6083 | -84.7711
5.6086 | -84.7713
5.609 | -84.772
5.6094 | -84.7726
5.6097 | -84.773
5.6101 | -84.7742
5.6104 | -84.777
5.6108 | -84.7779
5.6112 | -84.7789
5.6115 | -84.78
5.6119 | -84.7808
5.6123 | -84.781
5.6126 | -84.7813
5.613 | -84.7822
5.6133 | -84.783
5.6137 | -84.7834
5.6141 | -84.7836
5.6144 | -84.7845
5.6148 | -84.7851
5.6152 | -84.7865
5.6155 | -84.7873
5.6159 | -84.7884
5.6162 | -84.789
5.6166 | -84.789
5.617 | -84.7899
5.6173 | -84.7905
5.6177 | -84.7912
5.618 | -84.7925
5.6184 | -84.793
5.6188 | -84.793
5.6191 | -84.7935
5.6195 | -84.7947
5.6199 | -84.7977
5.6202 | -84.8007
5.6206 | -84.8013
5.6209 | -84.8023
5.6213 | -84.8033
5.6217 | -84.8046
5.622 | -84.8052
5.6224 | -84.806
5.6227 | -84.8059
5.6231 | -84.8058
5.6235 | -84.8062
5.6238 | -84.8067
5.6242 | -84.8071
5.6246 | -84.8094
5.6249 | -84.8111
5.6253 | -84.8119
5.6256 | -84.8125
5.626 | -84.8133
5.6264 | -84.8134
5.6267 | -84.8138
5.6271 | -84.8145
5.6275 | -84.8148
5.6278 | -84.8164
5.6282 | -84.8174
5.6285 | -84.8178
5.6289 | -84.8182
5.6293 | -84.8196
5.6296 | -84.8201
5.63 | -84.8198
5.6303 | -84.8205
5.6307 | -84.8207
5.6311 | -84.8205
5.6314 | -84.8206
5.6318 | -84.8211
5.6322 | -84.8227
5.6325 | -84.8239
5.6329 | -84.8245
5.6332 | -84.8249
5.6336 | -84.8255
5.634 | -84.8258
5.6343 | -84.8261
5.6347 | -84.8266
5.6351 | -84.8271
5.6354 | -84.8277
5.6358 | -84.8282
5.6361 | -84.8281
5.6365 | -84.827
5.6369 | -84.8269
5.6372 | -84.8274
5.6376 | -84.8283
5.6379 | -84.8289
5.6383 | -84.8291
5.6387 | -84.8297
5.639 | -84.8303
5.6394 | -84.8306
5.6398 | -84.8304
5.6401 | -84.83
5.6405 | -84.8308
5.6408 | -84.8306
5.6412 | -84.8302
5.6416 | -84.8309
5.6419 | -84.8311
5.6423 | -84.8318
5.6426 | -84.8317
5.643 | -84.8314
5.6434 | -84.831
5.6437 | -84.8303
5.6441 | -84.8316
5.6445 | -84.8324
5.6448 | -84.8329
5.6452 | -84.8332
5.6455 | -84.8334
5.6459 | -84.8337
5.6463 | -84.8335
5.6466 | -84.8341
5.647 | -84.8344
5.6474 | -84.8328
5.6477 | -84.8304
5.6481 | -84.8302
5.6484 | -84.8309
5.6488 | -84.8329
5.6492 | -84.8339
5.6495 | -84.8343
5.6499 | -84.8342
5.6502 | -84.8337
5.6506 | -84.8339
5.651 | -84.8345
5.6513 | -84.8342
5.6517 | -84.8347
5.6521 | -84.8345
5.6524 | -84.8346
5.6528 | -84.8347
5.6531 | -84.8359
5.6535 | -84.8378
5.6539 | -84.8396
5.6542 | -84.8409
5.6546 | -84.8411
5.655 | -84.8411
5.6553 | -84.8422
5.6557 | -84.8438
5.656 | -84.8451
5.6564 | -84.8449
5.6568 | -84.8442
5.6571 | -84.843
5.6575 | -84.8413
5.6578 | -84.8413
5.6582 | -84.841
5.6586 | -84.8406
5.6589 | -84.8399
5.6593 | -84.8393
5.6597 | -84.8385
5.66 | -84.8394
5.6604 | -84.84
5.6607 | -84.8397
5.6611 | -84.8396
5.6615 | -84.8396
5.6618 | -84.8392
5.6622 | -84.8379
5.6625 | -84.8374
5.6629 | -84.8372
5.6633 | -84.8378
5.6636 | -84.8384
5.664 | -84.8386
5.6644 | -84.8386
5.6647 | -84.8395
5.6651 | -84.8381
5.6654 | -84.8364
5.6658 | -84.8351
5.6662 | -84.8334
5.6665 | -84.8333
5.6669 | -84.8341
5.6673 | -84.8349
5.6676 | -84.8341
5.668 | -84.8324
5.6683 | -84.8322
5.6687 | -84.8328
5.6691 | -84.8334
5.6694 | -84.8332
5.6698 | -84.8333
5.6701 | -84.8324
5.6705 | -84.8314
5.6709 | -84.8311
5.6712 | -84.8308
5.6716 | -84.8301
5.672 | -84.8305
5.6723 | -84.8303
5.6727 | -84.8294
5.673 | -84.8279
5.6734 | -84.8275
5.6738 | -84.8269
5.6741 | -84.8271
5.6745 | -84.8281
5.6748 | -84.8281
5.6752 | -84.828
5.6756 | -84.8293
5.6759 | -84.8311
5.6763 | -84.8316
5.6767 | -84.8318
5.677 | -84.8309
5.6774 | -84.8303
5.6777 | -84.8296
5.6781 | -84.83
5.6785 | -84.8295
5.6788 | -84.8292
5.6792 | -84.8288
5.6796 | -84.8279
5.6799 | -84.8271
5.6803 | -84.8262
5.6806 | -84.8264
5.681 | -84.8265
5.6814 | -84.8247
5.6817 | -84.8238
5.6821 | -84.8233
5.6824 | -84.8228
5.6828 | -84.8227
5.6832 | -84.8237
5.6835 | -84.8243
5.6839 | -84.8252
5.6843 | -84.825
5.6846 | -84.8245
5.685 | -84.824
5.6853 | -84.8235
5.6857 | -84.8226
5.6861 | -84.8219
5.6864 | -84.8213
5.6868 | -84.8205
5.6872 | -84.8202
5.6875 | -84.8205
5.6879 | -84.8203
5.6882 | -84.8194
5.6886 | -84.8186
5.689 | -84.8185
5.6893 | -84.8182
5.6897 | -84.8176
5.69 | -84.8173
5.6904 | -84.8164
5.6908 | -84.816
5.6911 | -84.8161
5.6915 | -84.8157
5.6919 | -84.8147
5.6922 | -84.8124
5.6926 | -84.8114
5.6929 | -84.811
5.6933 | -84.8106
5.6937 | -84.8097
5.694 | -84.8095
5.6944 | -84.8093
5.6947 | -84.8091
5.6951 | -84.8071
5.6955 | -84.8057
5.6958 | -84.8052
5.6962 | -84.8056
5.6966 | -84.8063
5.6969 | -84.8058
5.6973 | -84.8054
5.6976 | -84.805
5.698 | -84.8042
5.6984 | -84.8036
5.6987 | -84.8029
5.6991 | -84.8024
5.6995 | -84.8019
5.6998 | -84.8007
5.7002 | -84.8
5.7005 | -84.7994
5.7009 | -84.7989
5.7013 | -84.7988
5.7016 | -84.7987
5.702 | -84.7983
5.7023 | -84.7991
5.7027 | -84.8003
5.7031 | -84.8
5.7034 | -84.7985
5.7038 | -84.7968
5.7042 | -84.7958
5.7045 | -84.7953
5.7049 | -84.795
5.7052 | -84.7939
5.7056 | -84.7934
5.706 | -84.7929
5.7063 | -84.7927
5.7067 | -84.7925
5.707 | -84.7926
5.7074 | -84.7931
5.7078 | -84.7932
5.7081 | -84.793
5.7085 | -84.7923
5.7089 | -84.7913
5.7092 | -84.7905
5.7096 | -84.7888
5.7099 | -84.7876
5.7103 | -84.7869
5.7107 | -84.7862
5.711 | -84.7856
5.7114 | -84.7848
5.7118 | -84.7837
5.7121 | -84.7829
5.7125 | -84.7826
5.7128 | -84.7821
5.7132 | -84.7809
5.7136 | -84.7789
5.7139 | -84.7785
5.7143 | -84.7778
5.7146 | -84.7771
5.715 | -84.776
5.7154 | -84.7753
5.7157 | -84.7745
5.7161 | -84.7738
5.7165 | -84.7729
5.7168 | -84.7726
5.7172 | -84.7732
5.7175 | -84.7727
5.7179 | -84.7716
5.7183 | -84.7708
5.7186 | -84.7703
5.719 | -84.7701
5.7194 | -84.7696
5.7197 | -84.7682
5.7201 | -84.7672
5.7204 | -84.7668
5.7208 | -84.7668
5.7212 | -84.7669
5.7215 | -84.7665
5.7219 | -84.7654
5.7222 | -84.7644
5.7226 | -84.7641
5.723 | -84.7634
5.7233 | -84.7624
5.7237 | -84.7619
5.7241 | -84.7614
5.7244 | -84.7607
5.7248 | -84.761
5.7251 | -84.76
5.7255 | -84.76
5.7259 | -84.7594
5.7262 | -84.7591
5.7266 | -84.7584
5.7269 | -84.7576
5.7273 | -84.7569
5.7277 | -84.756
5.728 | -84.7552
5.7284 | -84.7543
5.7288 | -84.7531
5.7291 | -84.7522
5.7295 | -84.7522
5.7298 | -84.7512
5.7302 | -84.7497
5.7306 | -84.7495
5.7309 | -84.7484
5.7313 | -84.7476
5.7317 | -84.7469
5.732 | -84.7461
5.7324 | -84.7453
5.7327 | -84.745
5.7331 | -84.7444
5.7335 | -84.7443
5.7338 | -84.7434
5.7342 | -84.7423
5.7345 | -84.7431
5.7349 | -84.7436
5.7353 | -84.7428
5.7356 | -84.7415
5.736 | -84.741
5.7364 | -84.7411
5.7367 | -84.741
5.7371 | -84.7411
5.7374 | -84.7406
5.7378 | -84.7393
5.7382 | -84.7373
5.7385 | -84.7358
5.7389 | -84.7355
5.7393 | -84.7349
5.7396 | -84.7349
5.74 | -84.7351
5.7403 | -84.7349
5.7407 | -84.7347
5.7411 | -84.7346
5.7414 | -84.733
5.7418 | -84.7312
5.7421 | -84.7302
5.7425 | -84.7296
5.7429 | -84.7283
5.7432 | -84.7284
5.7436 | -84.7281
5.744 | -84.7277
5.7443 | -84.7264
5.7447 | -84.7262
5.745 | -84.7256
5.7454 | -84.7253
5.7458 | -84.7246
5.7461 | -84.7231
5.7465 | -84.7206
5.7468 | -84.7191
5.7472 | -84.7188
5.7476 | -84.7192
5.7479 | -84.7198
5.7483 | -84.7187
5.7487 | -84.7182
5.749 | -84.7183
5.7494 | -84.718
5.7497 | -84.7168
5.7501 | -84.716
5.7505 | -84.7153
5.7508 | -84.7146
5.7512 | -84.7155
5.7516 | -84.7153
5.7519 | -84.718
5.7523 | -84.7182
5.7526 | -84.7172
5.753 | -84.7169
5.7534 | -84.716
5.7537 | -84.7156
5.7541 | -84.7148
5.7544 | -84.7144
5.7548 | -84.7143
5.7552 | -84.7141
5.7555 | -84.7127
5.7559 | -84.711
5.7563 | -84.71
5.7566 | -84.7093
5.757 | -84.7088
5.7573 | -84.7084
5.7577 | -84.7073
5.7581 | -84.7068
5.7584 | -84.7067
5.7588 | -84.7059
5.7591 | -84.7047
5.7595 | -84.7044
5.7599 | -84.704
5.7602 | -84.7027
5.7606 | -84.7021
5.761 | -84.7008
5.7613 | -84.7002
5.7617 | -84.6991
5.762 | -84.699
5.7624 | -84.698
5.7628 | -84.6976
5.7631 | -84.6968
5.7635 | -84.6962
5.7639 | -84.6957
5.7642 | -84.6959
5.7646 | -84.6957
5.7649 | -84.6953
5.7653 | -84.6949
5.7657 | -84.6942
5.766 | -84.6933
5.7664 | -84.6933
5.7667 | -84.6931
5.7671 | -84.694
5.7675 | -84.6934
5.7678 | -84.6912
5.7682 | -84.689
5.7686 | -84.6878
5.7689 | -84.6874
5.7693 | -84.6868
5.7696 | -84.6863
5.77 | -84.6841
5.7704 | -84.6823
5.7707 | -84.6818
5.7711 | -84.6814
5.7714 | -84.6807
5.7718 | -84.6808
5.7722 | -84.6815
5.7725 | -84.6815
5.7729 | -84.682
5.7733 | -84.6815
5.7736 | -84.681
5.774 | -84.6798
5.7743 | -84.679
5.7747 | -84.6784
5.7751 | -84.6779
5.7754 | -84.6779
5.7758 | -84.6769
5.7762 | -84.6765
5.7765 | -84.6762
5.7769 | -84.6756
5.7772 | -84.6746
5.7776 | -84.6743
5.778 | -84.6753
5.7783 | -84.6754
5.7787 | -84.675
5.779 | -84.6742
5.7794 | -84.6738
5.7798 | -84.6734
5.7801 | -84.6723
5.7805 | -84.6705
5.7809 | -84.6696
5.7812 | -84.6707
5.7816 | -84.6711
5.7819 | -84.6708
5.7823 | -84.6707
5.7827 | -84.6704
5.783 | -84.6711
5.7834 | -84.6707
5.7838 | -84.6703
5.7841 | -84.6697
5.7845 | -84.6695
5.7848 | -84.669
5.7852 | -84.669
5.7856 | -84.669
5.7859 | -84.6687
5.7863 | -84.6682
5.7866 | -84.6677
5.787 | -84.6673
5.7874 | -84.6671
5.7877 | -84.6668
5.7881 | -84.6676
5.7885 | -84.6682
5.7888 | -84.6674
5.7892 | -84.6666
5.7895 | -84.6664
5.7899 | -84.6655
5.7903 | -84.6649
5.7906 | -84.665
5.791 | -84.6659
5.7913 | -84.6657
5.7917 | -84.6658
5.7921 | -84.6657
5.7924 | -84.6653
5.7928 | -84.6637
5.7932 | -84.6628
5.7935 | -84.6626
5.7939 | -84.6625
5.7942 | -84.6613
5.7946 | -84.6606
5.795 | -84.6602
5.7953 | -84.6596
5.7957 | -84.66
5.7961 | -84.6597
5.7964 | -84.6594
5.7968 | -84.6594
5.7971 | -84.6596
5.7975 | -84.6599
5.7979 | -84.6594
5.7982 | -84.6589
5.7986 | -84.6596
5.7989 | -84.6598
5.7993 | -84.6592
5.7997 | -84.6585
5.8 | -84.6574
5.8004 | -84.6571
5.8008 | -84.6571
5.8011 | -84.657
5.8015 | -84.6561
5.8018 | -84.656
5.8022 | -84.6565
5.8026 | -84.657
5.8029 | -84.6558
5.8033 | -84.6536
5.8037 | -84.652
5.804 | -84.6512
5.8044 | -84.6511
5.8047 | -84.6495
5.8051 | -84.6479
5.8055 | -84.6468
5.8058 | -84.6465
5.8062 | -84.6471
5.8065 | -84.6469
5.8069 | -84.6468
5.8073 | -84.6473
5.8076 | -84.6477
5.808 | -84.6476
5.8084 | -84.6478
5.8087 | -84.6476
5.8091 | -84.6467
5.8094 | -84.646
5.8098 | -84.6445
5.8102 | -84.6428
5.8105 | -84.6423
5.8109 | -84.6424
5.8112 | -84.6424
5.8116 | -84.6424
5.812 | -84.6423
5.8123 | -84.6428
5.8127 | -84.6426
5.8131 | -84.6415
5.8134 | -84.6409
5.8138 | -84.6408
5.8141 | -84.6399
5.8145 | -84.6399
5.8149 | -84.6398
5.8152 | -84.6401
5.8156 | -84.6395
5.816 | -84.6392
5.8163 | -84.6393
5.8167 | -84.6387
5.817 | -84.6386
5.8174 | -84.6383
5.8178 | -84.6385
5.8181 | -84.6384
5.8185 | -84.6386
5.8188 | -84.6388
5.8192 | -84.6386
5.8196 | -84.6386
5.8199 | -84.6383
5.8203 | -84.6384
5.8207 | -84.6377
5.821 | -84.6375
5.8214 | -84.6375
5.8217 | -84.6369
5.8221 | -84.6356
5.8225 | -84.6353
5.8228 | -84.6346
5.8232 | -84.6332
5.8236 | -84.633
5.8239 | -84.6324
5.8243 | -84.631
5.8246 | -84.6297
5.825 | -84.6306
5.8254 | -84.6307
5.8257 | -84.6308
5.8261 | -84.6318
5.8264 | -84.6312
5.8268 | -84.6306
5.8272 | -84.63
5.8275 | -84.6303
5.8279 | -84.6306
5.8283 | -84.6303
5.8286 | -84.6306
5.829 | -84.6299
5.8293 | -84.6294
5.8297 | -84.6293
5.8301 | -84.6299
5.8304 | -84.6303
5.8308 | -84.6303
5.8311 | -84.6298
5.8315 | -84.6297
5.8319 | -84.6297
5.8322 | -84.6297
5.8326 | -84.6302
5.833 | -84.63
5.8333 | -84.6298
5.8337 | -84.6305
5.834 | -84.6308
5.8344 | -84.6306
5.8348 | -84.63
5.8351 | -84.6277
5.8355 | -84.6262
5.8359 | -84.6256
5.8362 | -84.6254
5.8366 | -84.6255
5.8369 | -84.6253
5.8373 | -84.6249
5.8377 | -84.6253
5.838 | -84.6274
5.8384 | -84.6281
5.8387 | -84.6277
5.8391 | -84.6275
5.8395 | -84.6277
5.8398 | -84.6279
5.8402 | -84.6279
5.8406 | -84.6275
5.8409 | -84.6281
5.8413 | -84.6279
5.8416 | -84.6271
5.842 | -84.6263
5.8424 | -84.6261
5.8427 | -84.626
5.8431 | -84.6264
5.8434 | -84.6266
5.8438 | -84.6268
5.8442 | -84.6265
5.8445 | -84.6261
5.8449 | -84.6267
5.8453 | -84.6266
5.8456 | -84.6276
5.846 | -84.6277
5.8463 | -84.6277
5.8467 | -84.6278
5.8471 | -84.6276
5.8474 | -84.6274
5.8478 | -84.6271
5.8482 | -84.6264
5.8485 | -84.6263
5.8489 | -84.6263
5.8492 | -84.6271
5.8496 | -84.6272
5.85 | -84.6273
5.8503 | -84.6267
5.8507 | -84.6257
5.851 | -84.6258
5.8514 | -84.6263
5.8518 | -84.6271
5.8521 | -84.6273
5.8525 | -84.6275
5.8529 | -84.6279
5.8532 | -84.628
5.8536 | -84.6281
5.8539 | -84.6281
5.8543 | -84.6282
5.8547 | -84.6285
5.855 | -84.6285
5.8554 | -84.6292
5.8557 | -84.6294
5.8561 | -84.6302
5.8565 | -84.6304
5.8568 | -84.6304
5.8572 | -84.6313
5.8576 | -84.6329
5.8579 | -84.6335
5.8583 | -84.6332
5.8586 | -84.6333
5.859 | -84.6334
5.8594 | -84.6338
5.8597 | -84.6342
5.8601 | -84.6344
5.8605 | -84.6346
5.8608 | -84.6348
5.8612 | -84.6349
5.8615 | -84.6349
5.8619 | -84.6349
5.8623 | -84.6352
5.8626 | -84.6361
5.863 | -84.6365
5.8633 | -84.6366
5.8637 | -84.6373
5.8641 | -84.6375
5.8644 | -84.6374
5.8648 | -84.6374
5.8652 | -84.6372
5.8655 | -84.6375
5.8659 | -84.6385
5.8662 | -84.6387
5.8666 | -84.6381
5.867 | -84.6396
5.8673 | -84.6409
5.8677 | -84.6408
5.8681 | -84.6403
5.8684 | -84.6407
5.8688 | -84.6416
5.8691 | -84.6422
5.8695 | -84.6429
5.8699 | -84.6426
5.8702 | -84.6429
5.8706 | -84.644
5.8709 | -84.6451
5.8713 | -84.6463
5.8717 | -84.6466
5.872 | -84.6471
5.8724 | -84.6477
5.8728 | -84.6484
5.8731 | -84.65
5.8735 | -84.6506
5.8738 | -84.6511
5.8742 | -84.6515
5.8746 | -84.6517
5.8749 | -84.6521
5.8753 | -84.6514
5.8756 | -84.6522
5.876 | -84.6523
5.8764 | -84.6526
5.8767 | -84.6524
5.8771 | -84.6524
5.8775 | -84.6527
5.8778 | -84.6538
5.8782 | -84.6544
5.8785 | -84.6545
5.8789 | -84.6555
5.8793 | -84.6562
5.8796 | -84.6557
5.88 | -84.6544
5.8804 | -84.6532
5.8807 | -84.6533
5.8811 | -84.6531
5.8814 | -84.6528
5.8818 | -84.6532
5.8822 | -84.653
5.8825 | -84.6526
5.8829 | -84.6516
5.8832 | -84.6528
5.8836 | -84.6528
5.884 | -84.653
5.8843 | -84.6536
5.8847 | -84.6536
5.8851 | -84.6536
5.8854 | -84.6539
5.8858 | -84.6543
5.8861 | -84.6554
5.8865 | -84.6559
5.8869 | -84.6562
5.8872 | -84.6563
5.8876 | -84.6568
5.888 | -84.6571
5.8883 | -84.6571
5.8887 | -84.6574
5.889 | -84.6577
5.8894 | -84.658
5.8898 | -84.6584
5.8901 | -84.6597
5.8905 | -84.6605
5.8908 | -84.6606
5.8912 | -84.6611
5.8916 | -84.6618
5.8919 | -84.6625
5.8923 | -84.6629
5.8927 | -84.6632
5.893 | -84.6633
5.8934 | -84.6636
5.8937 | -84.6644
5.8941 | -84.6652
5.8945 | -84.6656
5.8948 | -84.6661
5.8952 | -84.6664
5.8955 | -84.6663
5.8959 | -84.6671
5.8963 | -84.6677
5.8966 | -84.6681
5.897 | -84.6688
5.8974 | -84.6699
5.8977 | -84.6711
5.8981 | -84.6723
5.8984 | -84.673
5.8988 | -84.6738
5.8992 | -84.6743
5.8995 | -84.6744
5.8999 | -84.6738
5.9003 | -84.6738
5.9006 | -84.6735
5.901 | -84.6741
5.9013 | -84.6749
5.9017 | -84.6751
5.9021 | -84.6747
5.9024 | -84.6748
5.9028 | -84.6756
5.9031 | -84.6761
5.9035 | -84.6765
5.9039 | -84.6763
5.9042 | -84.6766
5.9046 | -84.6771
5.905 | -84.6774
5.9053 | -84.6776
5.9057 | -84.6776
5.906 | -84.6778
5.9064 | -84.6782
5.9068 | -84.6788
5.9071 | -84.6796
5.9075 | -84.6803
5.9078 | -84.6811
5.9082 | -84.6819
5.9086 | -84.6825
5.9089 | -84.6826
5.9093 | -84.6823
5.9097 | -84.6828
5.91 | -84.6845
5.9104 | -84.6852
5.9107 | -84.6859
5.9111 | -84.6868
5.9115 | -84.6877
5.9118 | -84.6891
5.9122 | -84.6898
5.9126 | -84.6902
5.9129 | -84.6909
5.9133 | -84.692
5.9136 | -84.6931
5.914 | -84.6938
5.9144 | -84.6944
5.9147 | -84.6948
5.9151 | -84.696
5.9154 | -84.6965
5.9158 | -84.6961
5.9162 | -84.6962
5.9165 | -84.6956
5.9169 | -84.6954
5.9173 | -84.6951
5.9176 | -84.6949
5.918 | -84.6949
5.9183 | -84.6956
5.9187 | -84.6963
5.9191 | -84.699
5.9194 | -84.7027
5.9198 | -84.7043
5.9202 | -84.704
5.9205 | -84.7043
5.9209 | -84.7046
5.9212 | -84.7042
5.9216 | -84.7037
5.922 | -84.7045
5.9223 | -84.7056
5.9227 | -84.7061
5.923 | -84.7065
5.9234 | -84.7079
5.9238 | -84.7085
5.9241 | -84.7091
5.9245 | -84.709
5.9249 | -84.7095
5.9252 | -84.7095
5.9256 | -84.7094
5.9259 | -84.7088
5.9263 | -84.7095
5.9267 | -84.7102
5.927 | -84.7112
5.9274 | -84.7118
5.9277 | -84.712
5.9281 | -84.712
5.9285 | -84.7125
5.9288 | -84.7135
5.9292 | -84.7125
5.9296 | -84.7134
5.9299 | -84.7141
5.9303 | -84.7144
5.9306 | -84.7148
5.931 | -84.7147
5.9314 | -84.7145
5.9317 | -84.7151
5.9321 | -84.7163
5.9325 | -84.7176
5.9328 | -84.7188
5.9332 | -84.7201
5.9335 | -84.721
5.9339 | -84.7221
5.9343 | -84.7227
5.9346 | -84.7229
5.935 | -84.7236
5.9353 | -84.7245
5.9357 | -84.7265
5.9361 | -84.7286
5.9364 | -84.729
5.9368 | -84.7292
5.9372 | -84.7305
5.9375 | -84.7317
5.9379 | -84.7322
5.9382 | -84.7328
5.9386 | -84.7337
5.939 | -84.7342
5.9393 | -84.7349
5.9397 | -84.7362
5.9401 | -84.737
5.9404 | -84.7375
5.9408 | -84.7376
5.9411 | -84.7381
5.9415 | -84.7388
5.9419 | -84.738
5.9422 | -84.7374
5.9426 | -84.7379
5.9429 | -84.7382
5.9433 | -84.7384
5.9437 | -84.738
5.944 | -84.7387
5.9444 | -84.7393
5.9448 | -84.7403
5.9451 | -84.7419
5.9455 | -84.7428
5.9458 | -84.7433
5.9462 | -84.7437
5.9466 | -84.7443
5.9469 | -84.7452
5.9473 | -84.7461
5.9476 | -84.7458
5.948 | -84.7461
5.9484 | -84.7467
5.9487 | -84.7473
5.9491 | -84.7476
5.9495 | -84.7485
5.9498 | -84.7499
5.9502 | -84.7504
5.9505 | -84.7507
5.9509 | -84.7514
5.9513 | -84.7518
5.9516 | -84.7524
5.952 | -84.753
5.9524 | -84.7529
5.9527 | -84.753
5.9531 | -84.754
5.9534 | -84.7548
5.9538 | -84.7556
5.9542 | -84.7559
5.9545 | -84.7555
5.9549 | -84.7561
5.9552 | -84.7569
5.9556 | -84.7577
5.956 | -84.7585
5.9563 | -84.7594
5.9567 | -84.7602
5.9571 | -84.761
5.9574 | -84.7615
5.9578 | -84.7621
5.9581 | -84.7632
5.9585 | -84.7646
5.9589 | -84.7656
5.9592 | -84.7662
5.9596 | -84.7669
5.9599 | -84.7679
5.9603 | -84.7676
5.9607 | -84.7686
5.961 | -84.77
5.9614 | -84.7707
5.9618 | -84.7713
5.9621 | -84.772
5.9625 | -84.7726
5.9628 | -84.7729
5.9632 | -84.7723
5.9636 | -84.7718
5.9639 | -84.771
5.9643 | -84.7712
5.9647 | -84.772
5.965 | -84.7729
5.9654 | -84.7737
5.9657 | -84.7736
5.9661 | -84.7742
5.9665 | -84.7754
5.9668 | -84.7761
5.9672 | -84.7769
5.9675 | -84.7774
5.9679 | -84.7786
5.9683 | -84.7797
5.9686 | -84.7813
5.969 | -84.7821
5.9694 | -84.7826
5.9697 | -84.7825
5.9701 | -84.7824
5.9704 | -84.7831
5.9708 | -84.7844
5.9712 | -84.7847
5.9715 | -84.7858
5.9719 | -84.7872
5.9722 | -84.788
5.9726 | -84.7903
5.973 | -84.7935
5.9733 | -84.7956
5.9737 | -84.7962
5.9741 | -84.7969
5.9744 | -84.7981
5.9748 | -84.7991
5.9751 | -84.8004
5.9755 | -84.8015
5.9759 | -84.8034
5.9762 | -84.805
5.9766 | -84.805
5.977 | -84.8058
5.9773 | -84.8067
5.9777 | -84.8073
5.978 | -84.8079
5.9784 | -84.8084
5.9788 | -84.8098
5.9791 | -84.8106
5.9795 | -84.8113
5.9798 | -84.8119
5.9802 | -84.8127
5.9806 | -84.8133
5.9809 | -84.8137
5.9813 | -84.8151
5.9817 | -84.8164
5.982 | -84.8173
5.9824 | -84.8183
5.9827 | -84.8192
5.9831 | -84.82
5.9835 | -84.8211
5.9838 | -84.8225
5.9842 | -84.8233
5.9846 | -84.8244
5.9849 | -84.8255
5.9853 | -84.8263
5.9856 | -84.827
5.986 | -84.8278
5.9864 | -84.829
5.9867 | -84.8303
5.9871 | -84.832
5.9874 | -84.8328
5.9878 | -84.8333
5.9882 | -84.8348
5.9885 | -84.8365
5.9889 | -84.8379
5.9893 | -84.8384
5.9896 | -84.839
5.99 | -84.8399
5.9903 | -84.841
5.9907 | -84.8431
5.9911 | -84.8449
5.9914 | -84.8459
5.9918 | -84.8478
5.9921 | -84.8487
5.9925 | -84.8501
5.9929 | -84.8518
5.9932 | -84.8525
5.9936 | -84.8536
5.994 | -84.8547
5.9943 | -84.8565
5.9947 | -84.8589
5.995 | -84.8611
5.9954 | -84.8619
5.9958 | -84.8621
5.9961 | -84.8628
5.9965 | -84.8637
5.9969 | -84.8642
5.9972 | -84.8649
5.9976 | -84.866
5.9979 | -84.8669
5.9983 | -84.868
5.9987 | -84.8696
5.999 | -84.8702
5.9994 | -84.8712
5.9997 | -84.8725
6.0001 | -84.874
6.0005 | -84.8752
6.0008 | -84.8772
6.0012 | -84.8792
6.0016 | -84.8797
6.0019 | -84.8798
6.0023 | -84.8815
6.0026 | -84.8822
6.003 | -84.8831
6.0034 | -84.8835
6.0037 | -84.8839
6.0041 | -84.8848
6.0045 | -84.8859
6.0048 | -84.8874
6.0052 | -84.8887
6.0055 | -84.8901
6.0059 | -84.8908
6.0063 | -84.8907
6.0066 | -84.8915
6.007 | -84.8923
6.0073 | -84.894
6.0077 | -84.8951
6.0081 | -84.8966
6.0084 | -84.8989
6.0088 | -84.9
6.0092 | -84.901
6.0095 | -84.902
6.0099 | -84.903
6.0102 | -84.9043
6.0106 | -84.9051
6.011 | -84.9064
6.0113 | -84.9082
6.0117 | -84.9117
6.012 | -84.9146
6.0124 | -84.916
6.0128 | -84.918
6.0131 | -84.9195
6.0135 | -84.921
6.0139 | -84.9223
6.0142 | -84.9234
6.0146 | -84.9241
6.0149 | -84.9246
6.0153 | -84.9255
6.0157 | -84.9267
6.016 | -84.9281
6.0164 | -84.9296
6.0168 | -84.9309
6.0171 | -84.9321
6.0175 | -84.9336
6.0178 | -84.9352
6.0182 | -84.9365
6.0186 | -84.9378
6.0189 | -84.9382
6.0193 | -84.9393
6.0196 | -84.9405
6.02 | -84.941
6.0204 | -84.9423
6.0207 | -84.944
6.0211 | -84.9451
6.0215 | -84.946
6.0218 | -84.9463
6.0222 | -84.9463
6.0225 | -84.9459
6.0229 | -84.9459
6.0233 | -84.9481
6.0236 | -84.9497
6.024 | -84.9507
6.0243 | -84.9521
6.0247 | -84.9536
6.0251 | -84.9551
6.0254 | -84.9559
6.0258 | -84.9575
6.0262 | -84.9584
6.0265 | -84.96
6.0269 | -84.9607
6.0272 | -84.9623
6.0276 | -84.964
6.028 | -84.9651
6.0283 | -84.9667
6.0287 | -84.9678
6.0291 | -84.969
6.0294 | -84.9699
6.0298 | -84.9707
6.0301 | -84.9718
6.0305 | -84.9736
6.0309 | -84.9749
6.0312 | -84.9743
6.0316 | -84.975
6.0319 | -84.9766
6.0323 | -84.978
6.0327 | -84.9787
6.033 | -84.9797
6.0334 | -84.9817
6.0338 | -84.9831
6.0341 | -84.9844
6.0345 | -84.9857
6.0348 | -84.9871
6.0352 | -84.9885
6.0356 | -84.9892
6.0359 | -84.99
6.0363 | -84.991
6.0366 | -84.9922
6.037 | -84.9937
6.0374 | -84.9942
6.0377 | -84.9933
6.0381 | -84.9935
6.0385 | -84.9945
6.0388 | -84.9954
6.0392 | -84.9973
6.0395 | -84.999
6.0399 | -84.9997
6.0403 | -85.0011
6.0406 | -85.0022
6.041 | -85.0034
6.0414 | -85.0054
6.0417 | -85.0082
6.0421 | -85.01
6.0424 | -85.0111
6.0428 | -85.012
6.0432 | -85.0132
6.0435 | -85.0134
6.0439 | -85.0144
6.0442 | -85.0147
6.0446 | -85.0154
6.045 | -85.0156
6.0453 | -85.0166
6.0457 | -85.0187
6.0461 | -85.021
6.0464 | -85.0223
6.0468 | -85.0241
6.0471 | -85.026
6.0475 | -85.0272
6.0479 | -85.0287
6.0482 | -85.0289
6.0486 | -85.0301
6.049 | -85.0314
6.0493 | -85.0329
6.0497 | -85.0353
6.05 | -85.0363
6.0504 | -85.0379
6.0508 | -85.0387
6.0511 | -85.0397
6.0515 | -85.0413
6.0518 | -85.043
6.0522 | -85.0451
6.0526 | -85.0461
6.0529 | -85.0471
6.0533 | -85.0481
6.0537 | -85.05
6.054 | -85.0524
6.0544 | -85.0539
6.0547 | -85.0557
6.0551 | -85.057
6.0555 | -85.0574
6.0558 | -85.0581
6.0562 | -85.0592
6.0565 | -85.0601
6.0569 | -85.061
6.0573 | -85.0627
6.0576 | -85.0641
6.058 | -85.0654
6.0584 | -85.0661
6.0587 | -85.0679
6.0591 | -85.0692
6.0594 | -85.0697
6.0598 | -85.07
6.0602 | -85.0707
6.0605 | -85.0716
6.0609 | -85.0727
6.0613 | -85.074
6.0616 | -85.0754
6.062 | -85.0766
6.0623 | -85.0772
6.0627 | -85.0765
6.0631 | -85.0768
6.0634 | -85.0774
6.0638 | -85.0794
6.0641 | -85.082
6.0645 | -85.0827
6.0649 | -85.0826
6.0652 | -85.0832
6.0656 | -85.0843
6.066 | -85.0855
6.0663 | -85.0863
6.0667 | -85.0882
6.067 | -85.0903
6.0674 | -85.0913
6.0678 | -85.0929
6.0681 | -85.0948
6.0685 | -85.0959
6.0689 | -85.0971
6.0692 | -85.0985
6.0696 | -85.1005
6.0699 | -85.102
6.0703 | -85.1028
6.0707 | -85.1038
6.071 | -85.1053
6.0714 | -85.1063
6.0717 | -85.1073
6.0721 | -85.1084
6.0725 | -85.1109
6.0728 | -85.1134
6.0732 | -85.1156
6.0736 | -85.1182
6.0739 | -85.1192
6.0743 | -85.1202
6.0746 | -85.1209
6.075 | -85.1221
6.0754 | -85.1238
6.0757 | -85.126
6.0761 | -85.1277
6.0764 | -85.1287
6.0768 | -85.1295
6.0772 | -85.1309
6.0775 | -85.1322
6.0779 | -85.1328
6.0783 | -85.1345
6.0786 | -85.1368
6.079 | -85.1377
6.0793 | -85.1379
6.0797 | -85.1393
6.0801 | -85.1403
6.0804 | -85.1414
6.0808 | -85.1428
6.0812 | -85.1447
6.0815 | -85.1451
6.0819 | -85.1453
6.0822 | -85.1454
6.0826 | -85.1462
6.083 | -85.1471
6.0833 | -85.148
6.0837 | -85.1489
6.084 | -85.1499
6.0844 | -85.151
6.0848 | -85.1517
6.0851 | -85.1531
6.0855 | -85.1532
6.0859 | -85.1543
6.0862 | -85.1561
6.0866 | -85.1574
6.0869 | -85.1584
6.0873 | -85.1598
6.0877 | -85.1608
6.088 | -85.1617
6.0884 | -85.1627
6.0888 | -85.164
6.0891 | -85.1657
6.0895 | -85.167
6.0898 | -85.1683
6.0902 | -85.1708
6.0906 | -85.172
6.0909 | -85.1728
6.0913 | -85.1734
6.0916 | -85.1746
6.092 | -85.1758
6.0924 | -85.1768
6.0927 | -85.1777
6.0931 | -85.1786
6.0935 | -85.1794
6.0938 | -85.1806
6.0942 | -85.1818
6.0945 | -85.183
6.0949 | -85.183
6.0953 | -85.1836
6.0956 | -85.1848
6.096 | -85.1852
6.0963 | -85.1865
6.0967 | -85.1876
6.0971 | -85.1882
6.0974 | -85.1888
6.0978 | -85.1899
6.0982 | -85.191
6.0985 | -85.1921
6.0989 | -85.193
6.0992 | -85.1934
6.0996 | -85.1933
6.1 | -85.1942
6.1003 | -85.1955
6.1007 | -85.196
6.1011 | -85.1963
6.1014 | -85.1974
6.1018 | -85.1998
6.1021 | -85.2004
6.1025 | -85.2007
6.1029 | -85.2015
6.1032 | -85.2024
6.1036 | -85.2029
6.1039 | -85.2029
6.1043 | -85.2034
6.1047 | -85.2056
6.105 | -85.2073
6.1054 | -85.2086
6.1058 | -85.2095
6.1061 | -85.2099
6.1065 | -85.2104
6.1068 | -85.2112
6.1072 | -85.2123
6.1076 | -85.2133
6.1079 | -85.2143
6.1083 | -85.2151
6.1086 | -85.2159
6.109 | -85.217
6.1094 | -85.2179
6.1097 | -85.2191
6.1101 | -85.2203
6.1105 | -85.2212
6.1108 | -85.2225
6.1112 | -85.2237
6.1115 | -85.2244
6.1119 | -85.2253
6.1123 | -85.2264
6.1126 | -85.2272
6.113 | -85.2284
6.1134 | -85.2294
6.1137 | -85.2313
6.1141 | -85.2332
6.1144 | -85.2343
6.1148 | -85.2356
6.1152 | -85.2365
6.1155 | -85.2381
6.1159 | -85.2384
6.1162 | -85.2394
6.1166 | -85.2396
6.117 | -85.2397
6.1173 | -85.2411
6.1177 | -85.2419
6.1181 | -85.2427
6.1184 | -85.2431
6.1188 | -85.2437
6.1191 | -85.2445
6.1195 | -85.2457
6.1199 | -85.2462
6.1202 | -85.2475
6.1206 | -85.2491
6.1209 | -85.2494
6.1213 | -85.2483
6.1217 | -85.2485
6.122 | -85.2493
6.1224 | -85.2503
6.1228 | -85.2515
6.1231 | -85.2521
6.1235 | -85.2518
6.1238 | -85.2517
6.1242 | -85.2535
6.1246 | -85.2547
6.1249 | -85.255
6.1253 | -85.2554
6.1257 | -85.2563
6.126 | -85.2575
6.1264 | -85.2579
6.1267 | -85.2585
6.1271 | -85.2592
6.1275 | -85.2598
6.1278 | -85.2604
6.1282 | -85.2615
6.1285 | -85.2625
6.1289 | -85.2636
6.1293 | -85.2649
6.1296 | -85.2658
6.13 | -85.2666
6.1304 | -85.2676
6.1307 | -85.269
6.1311 | -85.2704
6.1314 | -85.2707
6.1318 | -85.2711
6.1322 | -85.2718
6.1325 | -85.2722
6.1329 | -85.2728
6.1333 | -85.2727
6.1336 | -85.2729
6.134 | -85.2734
6.1343 | -85.2739
6.1347 | -85.2754
6.1351 | -85.2764
6.1354 | -85.2752
6.1358 | -85.2756
6.1361 | -85.2763
6.1365 | -85.277
6.1369 | -85.2783
6.1372 | -85.2788
6.1376 | -85.2799
6.138 | -85.2808
6.1383 | -85.2816
6.1387 | -85.2818
6.139 | -85.2821
6.1394 | -85.2831
6.1398 | -85.2847
6.1401 | -85.2857
6.1405 | -85.2862
6.1408 | -85.2866
6.1412 | -85.2881
6.1416 | -85.2889
6.1419 | -85.29
6.1423 | -85.2909
6.1427 | -85.2916
6.143 | -85.2923
6.1434 | -85.2933
6.1437 | -85.2942
6.1441 | -85.295
6.1445 | -85.2961
6.1448 | -85.2968
6.1452 | -85.2976
6.1456 | -85.2987
6.1459 | -85.2988
6.1463 | -85.2993
6.1466 | -85.3004
6.147 | -85.3012
6.1474 | -85.3014
6.1477 | -85.3017
6.1481 | -85.3032
6.1484 | -85.3033
6.1488 | -85.3039
6.1492 | -85.3043
6.1495 | -85.3046
6.1499 | -85.3053
6.1503 | -85.3057
6.1506 | -85.3067
6.151 | -85.3089
6.1513 | -85.3087
6.1517 | -85.3079
6.1521 | -85.3079
6.1524 | -85.3078
6.1528 | -85.3079
6.1532 | -85.3082
6.1535 | -85.3086
6.1539 | -85.3082
6.1542 | -85.3086
6.1546 | -85.3106
6.155 | -85.3115
6.1553 | -85.3105
6.1557 | -85.3103
6.156 | -85.311
6.1564 | -85.3115
6.1568 | -85.3115
6.1571 | -85.3122
6.1575 | -85.3124
6.1579 | -85.3125
6.1582 | -85.313
6.1586 | -85.3134
6.1589 | -85.3137
6.1593 | -85.3133
6.1597 | -85.3134
6.16 | -85.3155
6.1604 | -85.3168
6.1607 | -85.3172
6.1611 | -85.3176
6.1615 | -85.3179
6.1618 | -85.3198
6.1622 | -85.3219
6.1626 | -85.3232
6.1629 | -85.3226
6.1633 | -85.3235
6.1636 | -85.3239
6.164 | -85.3237
6.1644 | -85.3237
6.1647 | -85.3237
6.1651 | -85.324
6.1655 | -85.3248
6.1658 | -85.3259
6.1662 | -85.3265
6.1665 | -85.3273
6.1669 | -85.3278
6.1673 | -85.3281
6.1676 | -85.3284
6.168 | -85.3287
6.1683 | -85.3288
6.1687 | -85.3289
6.1691 | -85.3296
6.1694 | -85.3298
6.1698 | -85.3312
6.1702 | -85.3329
6.1705 | -85.3337
6.1709 | -85.3341
6.1712 | -85.3334
6.1716 | -85.3328
6.172 | -85.3335
6.1723 | -85.3335
6.1727 | -85.3335
6.173 | -85.3339
6.1734 | -85.3351
6.1738 | -85.336
6.1741 | -85.3364
6.1745 | -85.3367
6.1749 | -85.336
6.1752 | -85.3367
6.1756 | -85.3369
6.1759 | -85.3375
6.1763 | -85.3379
6.1767 | -85.3384
6.177 | -85.3391
6.1774 | -85.3393
6.1778 | -85.3402
6.1781 | -85.3405
6.1785 | -85.3396
6.1788 | -85.3398
6.1792 | -85.3397
6.1796 | -85.3401
6.1799 | -85.3402
6.1803 | -85.341
6.1806 | -85.3418
6.181 | -85.3428
6.1814 | -85.3431
6.1817 | -85.3442
6.1821 | -85.3447
6.1825 | -85.3457
6.1828 | -85.3477
6.1832 | -85.3495
6.1835 | -85.3503
6.1839 | -85.3498
6.1843 | -85.3498
6.1846 | -85.3505
6.185 | -85.3523
6.1854 | -85.3536
6.1857 | -85.3539
6.1861 | -85.3539
6.1864 | -85.3538
6.1868 | -85.3543
6.1872 | -85.3548
6.1875 | -85.3551
6.1879 | -85.3556
6.1882 | -85.3565
6.1886 | -85.3573
6.189 | -85.3573
6.1893 | -85.3579
6.1897 | -85.3579
6.1901 | -85.3583
6.1904 | -85.3583
6.1908 | -85.3582
6.1911 | -85.3584
6.1915 | -85.3589
6.1919 | -85.3591
6.1922 | -85.3591
6.1926 | -85.3588
6.1929 | -85.3598
6.1933 | -85.3603
6.1937 | -85.3605
6.194 | -85.3606
6.1944 | -85.3611
6.1948 | -85.3617
6.1951 | -85.3623
6.1955 | -85.363
6.1958 | -85.364
6.1962 | -85.3656
6.1966 | -85.3666
6.1969 | -85.3667
6.1973 | -85.3665
6.1977 | -85.3669
6.198 | -85.3672
6.1984 | -85.3674
6.1987 | -85.3673
6.1991 | -85.3668
6.1995 | -85.3667
6.1998 | -85.3668
6.2002 | -85.3671
6.2005 | -85.3682
6.2009 | -85.3684
6.2013 | -85.369
6.2016 | -85.3684
6.202 | -85.3684
6.2024 | -85.3688
6.2027 | -85.3692
6.2031 | -85.3691
6.2034 | -85.3692
6.2038 | -85.3696
6.2042 | -85.37
6.2045 | -85.3702
6.2049 | -85.3711
6.2053 | -85.3723
6.2056 | -85.373
6.206 | -85.3735
6.2063 | -85.3745
6.2067 | -85.375
6.2071 | -85.3749
6.2074 | -85.375
6.2078 | -85.3755
6.2081 | -85.3755
6.2085 | -85.3751
6.2089 | -85.3757
6.2092 | -85.3776
6.2096 | -85.3784
6.21 | -85.3788
6.2103 | -85.3793
6.2107 | -85.3802
6.211 | -85.3811
6.2114 | -85.3821
6.2118 | -85.3824
6.2121 | -85.3825
6.2125 | -85.3825
6.2128 | -85.3823
6.2132 | -85.3832
6.2136 | -85.3835
6.2139 | -85.384
6.2143 | -85.3844
6.2147 | -85.3847
6.215 | -85.3852
6.2154 | -85.3856
6.2157 | -85.3856
6.2161 | -85.3848
6.2165 | -85.3839
6.2168 | -85.3838
6.2172 | -85.3845
6.2176 | -85.3847
6.2179 | -85.3848
6.2183 | -85.3847
6.2186 | -85.3859
6.219 | -85.3862
6.2194 | -85.3866
6.2197 | -85.3866
6.2201 | -85.3879
6.2204 | -85.3906
6.2208 | -85.3923
6.2212 | -85.3931
6.2215 | -85.3942
6.2219 | -85.3952
6.2223 | -85.3956
6.2226 | -85.3974
6.223 | -85.399
6.2233 | -85.4002
6.2237 | -85.4009
6.2241 | -85.401
6.2244 | -85.4013
6.2248 | -85.4015
6.2251 | -85.4018
6.2255 | -85.4021
6.2259 | -85.4026
6.2262 | -85.4032
6.2266 | -85.4036
6.227 | -85.4046
6.2273 | -85.4056
6.2277 | -85.4075
6.228 | -85.4094
6.2284 | -85.41
6.2288 | -85.4103
6.2291 | -85.4105
6.2295 | -85.411
6.2299 | -85.4118
6.2302 | -85.4121
6.2306 | -85.4124
6.2309 | -85.4135
6.2313 | -85.4139
6.2317 | -85.4149
6.232 | -85.4153
6.2324 | -85.4168
6.2327 | -85.4178
6.2331 | -85.4185
6.2335 | -85.4195
6.2338 | -85.4193
6.2342 | -85.4195
6.2346 | -85.4201
6.2349 | -85.4205
6.2353 | -85.4217
6.2356 | -85.4222
6.236 | -85.4215
6.2364 | -85.4198
6.2367 | -85.4184
6.2371 | -85.4175
6.2374 | -85.4178
6.2378 | -85.4182
6.2382 | -85.4185
6.2385 | -85.4191
6.2389 | -85.4194
6.2393 | -85.4203
6.2396 | -85.4209
6.24 | -85.4213
6.2403 | -85.4225
6.2407 | -85.4247
6.2411 | -85.4253
6.2414 | -85.4261
6.2418 | -85.4277
6.2422 | -85.4279
6.2425 | -85.4293
6.2429 | -85.4308
6.2432 | -85.4313
6.2436 | -85.4319
6.244 | -85.4323
6.2443 | -85.4323
6.2447 | -85.4326
6.245 | -85.4327
6.2454 | -85.4332
6.2458 | -85.4337
6.2461 | -85.4339
6.2465 | -85.4348
6.2469 | -85.4352
6.2472 | -85.4358
6.2476 | -85.4369
6.2479 | -85.4382
6.2483 | -85.4393
6.2487 | -85.4399
6.249 | -85.4404
6.2494 | -85.4409
6.2498 | -85.4414
6.2501 | -85.4419
6.2505 | -85.4427
6.2508 | -85.443
6.2512 | -85.4431
6.2516 | -85.4429
6.2519 | -85.4426
6.2523 | -85.4427
6.2526 | -85.4436
6.253 | -85.444
6.2534 | -85.444
6.2537 | -85.4445
6.2541 | -85.4447
6.2545 | -85.4454
6.2548 | -85.4474
6.2552 | -85.4499
6.2555 | -85.452
6.2559 | -85.4528
6.2563 | -85.4534
6.2566 | -85.4538
6.257 | -85.4545
6.2573 | -85.4551
6.2577 | -85.456
6.2581 | -85.4565
6.2584 | -85.4571
6.2588 | -85.458
6.2592 | -85.4583
6.2595 | -85.4583
6.2599 | -85.4588
6.2602 | -85.4591
6.2606 | -85.4598
6.261 | -85.4604
6.2613 | -85.4607
6.2617 | -85.4617
6.2621 | -85.4629
6.2624 | -85.4638
6.2628 | -85.4645
6.2631 | -85.4658
6.2635 | -85.4659
6.2639 | -85.4665
6.2642 | -85.4674
6.2646 | -85.4681
6.2649 | -85.4695
6.2653 | -85.4701
6.2657 | -85.4708
6.266 | -85.4712
6.2664 | -85.4714
6.2668 | -85.472
6.2671 | -85.4723
6.2675 | -85.4726
6.2678 | -85.4728
6.2682 | -85.4735
6.2686 | -85.4747
6.2689 | -85.4761
6.2693 | -85.4769
6.2697 | -85.4775
6.27 | -85.4783
6.2704 | -85.4791
6.2707 | -85.4793
6.2711 | -85.4805
6.2715 | -85.4823
6.2718 | -85.4825
6.2722 | -85.4825
6.2725 | -85.4825
6.2729 | -85.4826
6.2733 | -85.4828
6.2736 | -85.4844
6.274 | -85.485
6.2744 | -85.4851
6.2747 | -85.4856
6.2751 | -85.4862
6.2754 | -85.4876
6.2758 | -85.4882
6.2762 | -85.489
6.2765 | -85.4895
6.2769 | -85.4895
6.2772 | -85.4903
6.2776 | -85.4911
6.278 | -85.4919
6.2783 | -85.4927
6.2787 | -85.4937
6.2791 | -85.4946
6.2794 | -85.4954
6.2798 | -85.4962
6.2801 | -85.4971
6.2805 | -85.4984
6.2809 | -85.4991
6.2812 | -85.4992
6.2816 | -85.4996
6.282 | -85.5003
6.2823 | -85.501
6.2827 | -85.5021
6.283 | -85.5022
6.2834 | -85.5017
6.2838 | -85.5018
6.2841 | -85.5024
6.2845 | -85.5033
6.2848 | -85.5039
6.2852 | -85.505
6.2856 | -85.5058
6.2859 | -85.5068
6.2863 | -85.508
6.2867 | -85.5092
6.287 | -85.5108
6.2874 | -85.5122
6.2877 | -85.5131
6.2881 | -85.5137
6.2885 | -85.516
6.2888 | -85.5173
6.2892 | -85.5177
6.2895 | -85.5188
6.2899 | -85.5206
6.2903 | -85.5223
6.2906 | -85.5236
6.291 | -85.5245
6.2914 | -85.5254
6.2917 | -85.526
6.2921 | -85.5274
6.2924 | -85.5283
6.2928 | -85.5284
6.2932 | -85.5286
6.2935 | -85.5287
6.2939 | -85.529
6.2943 | -85.5295
6.2946 | -85.5312
6.295 | -85.532
6.2953 | -85.5328
6.2957 | -85.534
6.2961 | -85.5347
6.2964 | -85.5355
6.2968 | -85.5368
6.2971 | -85.5378
6.2975 | -85.5394
6.2979 | -85.5412
6.2982 | -85.542
6.2986 | -85.5442
6.299 | -85.5457
6.2993 | -85.5477
6.2997 | -85.5486
6.3 | -85.5497
6.3004 | -85.5511
6.3008 | -85.5521
6.3011 | -85.5534
6.3015 | -85.5547
6.3019 | -85.5557
6.3022 | -85.5568
6.3026 | -85.5586
6.3029 | -85.5614
6.3033 | -85.5641
6.3037 | -85.5654
6.304 | -85.567
6.3044 | -85.5688
6.3047 | -85.5708
6.3051 | -85.5719
6.3055 | -85.5735
6.3058 | -85.5752
6.3062 | -85.5778
6.3066 | -85.5789
6.3069 | -85.58
6.3073 | -85.5803
6.3076 | -85.5814
6.308 | -85.5831
6.3084 | -85.5848
6.3087 | -85.5861
6.3091 | -85.5872
6.3094 | -85.5878
6.3098 | -85.5877
6.3102 | -85.588
6.3105 | -85.5891
6.3109 | -85.5909
6.3113 | -85.5918
6.3116 | -85.593
6.312 | -85.5946
6.3123 | -85.5957
6.3127 | -85.5961
6.3131 | -85.596
6.3134 | -85.597
6.3138 | -85.5979
6.3142 | -85.5987
6.3145 | -85.6
6.3149 | -85.6005
6.3152 | -85.6022
6.3156 | -85.6033
6.316 | -85.6047
6.3163 | -85.606
6.3167 | -85.6072
6.317 | -85.6086
6.3174 | -85.6101
6.3178 | -85.6114
6.3181 | -85.6127
6.3185 | -85.6139
6.3189 | -85.6148
6.3192 | -85.6157
6.3196 | -85.617
6.3199 | -85.6181
6.3203 | -85.6194
6.3207 | -85.6198
6.321 | -85.6205
6.3214 | -85.6212
6.3217 | -85.6223
6.3221 | -85.6236
6.3225 | -85.6249
6.3228 | -85.6264
6.3232 | -85.6276
6.3236 | -85.6296
6.3239 | -85.6319
6.3243 | -85.6344
6.3246 | -85.6351
6.325 | -85.6365
6.3254 | -85.6377
6.3257 | -85.6396
6.3261 | -85.6401
6.3265 | -85.6417
6.3268 | -85.644
6.3272 | -85.6454
6.3275 | -85.647
6.3279 | -85.6481
6.3283 | -85.6492
6.3286 | -85.65
6.329 | -85.6506
6.3293 | -85.6519
6.3297 | -85.653
6.3301 | -85.6539
6.3304 | -85.6557
6.3308 | -85.6573
6.3312 | -85.6575
6.3315 | -85.6583
6.3319 | -85.66
6.3322 | -85.6616
6.3326 | -85.6627
6.333 | -85.6632
6.3333 | -85.6649
6.3337 | -85.6664
6.3341 | -85.6681
6.3344 | -85.6702
6.3348 | -85.6717
6.3351 | -85.673
6.3355 | -85.6742
6.3359 | -85.6759
6.3362 | -85.6775
6.3366 | -85.6794
6.3369 | -85.6806
6.3373 | -85.6817
6.3377 | -85.6826
6.338 | -85.6816
6.3384 | -85.6822
6.3388 | -85.6834
6.3391 | -85.6841
6.3395 | -85.6848
6.3398 | -85.6833
6.3402 | -85.6825
6.3406 | -85.6836
6.3409 | -85.6851
6.3413 | -85.6874
6.3416 | -85.6893
6.342 | -85.6914
6.3424 | -85.6926
6.3427 | -85.6946
6.3431 | -85.6963
6.3435 | -85.6982
6.3438 | -85.6994
6.3442 | -85.7012
6.3445 | -85.7026
6.3449 | -85.7039
6.3453 | -85.7049
6.3456 | -85.7058
6.346 | -85.7074
6.3464 | -85.7088
6.3467 | -85.7105
6.3471 | -85.7111
6.3474 | -85.7128
6.3478 | -85.715
6.3482 | -85.7156
6.3485 | -85.7162
6.3489 | -85.7171
6.3492 | -85.717
6.3496 | -85.718
6.35 | -85.7192
6.3503 | -85.7205
6.3507 | -85.7225
6.3511 | -85.724
6.3514 | -85.7254
6.3518 | -85.727
6.3521 | -85.7292
6.3525 | -85.7307
6.3529 | -85.7316
6.3532 | -85.7324
6.3536 | -85.7327
6.354 | -85.7335
6.3543 | -85.7353
6.3547 | -85.7368
6.355 | -85.738
6.3554 | -85.7398
6.3558 | -85.7429
6.3561 | -85.7444
6.3565 | -85.7457
6.3568 | -85.7469
6.3572 | -85.747
6.3576 | -85.7476
6.3579 | -85.7494
6.3583 | -85.7507
6.3587 | -85.7511
6.359 | -85.7514
6.3594 | -85.7527
6.3597 | -85.7538
6.3601 | -85.755
6.3605 | -85.7558
6.3608 | -85.7563
6.3612 | -85.757
6.3615 | -85.7579
6.3619 | -85.7583
6.3623 | -85.7591
6.3626 | -85.7594
6.363 | -85.76
6.3634 | -85.7618
6.3637 | -85.7638
6.3641 | -85.7651
6.3644 | -85.7661
6.3648 | -85.7673
6.3652 | -85.7679
6.3655 | -85.7685
6.3659 | -85.7694
6.3663 | -85.7705
6.3666 | -85.7711
6.367 | -85.7712
6.3673 | -85.7721
6.3677 | -85.7733
6.3681 | -85.7746
6.3684 | -85.7765
6.3688 | -85.7782
6.3691 | -85.7796
6.3695 | -85.7825
6.3699 | -85.7843
6.3702 | -85.7854
6.3706 | -85.7867
6.371 | -85.7885
6.3713 | -85.7907
6.3717 | -85.7919
6.372 | -85.7929
6.3724 | -85.7935
6.3728 | -85.7944
6.3731 | -85.7955
6.3735 | -85.7967
6.3738 | -85.798
6.3742 | -85.7991
6.3746 | -85.8007
6.3749 | -85.8029
6.3753 | -85.805
6.3757 | -85.8065
6.376 | -85.8075
6.3764 | -85.8088
6.3767 | -85.8104
6.3771 | -85.8113
6.3775 | -85.8123
6.3778 | -85.8137
6.3782 | -85.8149
6.3786 | -85.816
6.3789 | -85.8175
6.3793 | -85.8188
6.3796 | -85.8202
6.38 | -85.8214
6.3804 | -85.8224
6.3807 | -85.8245
6.3811 | -85.8265
6.3814 | -85.8274
6.3818 | -85.8285
6.3822 | -85.8296
6.3825 | -85.8311
6.3829 | -85.8325
6.3833 | -85.8331
6.3836 | -85.8342
6.384 | -85.8349
6.3843 | -85.8352
6.3847 | -85.8355
6.3851 | -85.8364
6.3854 | -85.8377
6.3858 | -85.8395
6.3861 | -85.8428
6.3865 | -85.8439
6.3869 | -85.8447
6.3872 | -85.8459
6.3876 | -85.847
6.388 | -85.8482
6.3883 | -85.8492
6.3887 | -85.8495
6.389 | -85.8513
6.3894 | -85.8528
6.3898 | -85.8529
6.3901 | -85.8536
6.3905 | -85.855
6.3909 | -85.8563
6.3912 | -85.8574
6.3916 | -85.8582
6.3919 | -85.8594
6.3923 | -85.862
6.3927 | -85.8641
6.393 | -85.8667
6.3934 | -85.8676
6.3937 | -85.8674
6.3941 | -85.8677
6.3945 | -85.8676
6.3948 | -85.8686
6.3952 | -85.8701
6.3956 | -85.8716
6.3959 | -85.8722
6.3963 | -85.8731
6.3966 | -85.8742
6.397 | -85.8751
6.3974 | -85.8759
6.3977 | -85.8768
6.3981 | -85.8778
6.3985 | -85.8792
6.3988 | -85.8803
6.3992 | -85.8825
6.3995 | -85.8846
6.3999 | -85.8861
6.4003 | -85.8866
6.4006 | -85.8867
6.401 | -85.8866
6.4013 | -85.8857
6.4017 | -85.8856
6.4021 | -85.8863
6.4024 | -85.8868
6.4028 | -85.8871
6.4032 | -85.8876
6.4035 | -85.8886
6.4039 | -85.8879
6.4042 | -85.8867
6.4046 | -85.8867
6.405 | -85.8873
6.4053 | -85.8878
6.4057 | -85.8888
6.406 | -85.8902
6.4064 | -85.8918
6.4068 | -85.893
6.4071 | -85.8941
6.4075 | -85.8951
6.4079 | -85.8961
6.4082 | -85.8963
6.4086 | -85.8962
6.4089 | -85.8961
6.4093 | -85.8959
6.4097 | -85.8965
6.41 | -85.8972
6.4104 | -85.8973
6.4108 | -85.8971
6.4111 | -85.8981
6.4115 | -85.8994
6.4118 | -85.9001
6.4122 | -85.901
6.4126 | -85.9019
6.4129 | -85.9024
6.4133 | -85.9033
6.4136 | -85.9041
6.414 | -85.9054
6.4144 | -85.9062
6.4147 | -85.9064
6.4151 | -85.9068
6.4155 | -85.9074
6.4158 | -85.9078
6.4162 | -85.9085
6.4165 | -85.9093
6.4169 | -85.9099
6.4173 | -85.9116
6.4176 | -85.9123
6.418 | -85.9133
6.4184 | -85.9144
6.4187 | -85.9151
6.4191 | -85.9154
6.4194 | -85.9147
6.4198 | -85.9151
6.4202 | -85.9162
6.4205 | -85.9171
6.4209 | -85.917
6.4212 | -85.9176
6.4216 | -85.918
6.422 | -85.9182
6.4223 | -85.9187
6.4227 | -85.9192
6.4231 | -85.92
6.4234 | -85.9192
6.4238 | -85.9194
6.4241 | -85.9201
6.4245 | -85.921
6.4249 | -85.9208
6.4252 | -85.9203
6.4256 | -85.9202
6.4259 | -85.9207
6.4263 | -85.9206
6.4267 | -85.9221
6.427 | -85.9233
6.4274 | -85.9236
6.4278 | -85.9243
6.4281 | -85.9257
6.4285 | -85.9251
6.4288 | -85.9242
6.4292 | -85.9241
6.4296 | -85.924
6.4299 | -85.9241
6.4303 | -85.9243
6.4307 | -85.9251
6.431 | -85.9261
6.4314 | -85.9273
6.4317 | -85.9285
6.4321 | -85.9292
6.4325 | -85.9296
6.4328 | -85.931
6.4332 | -85.9321
6.4335 | -85.9336
6.4339 | -85.9344
6.4343 | -85.9346
6.4346 | -85.935
6.435 | -85.9354
6.4354 | -85.9365
6.4357 | -85.9373
6.4361 | -85.9369
6.4364 | -85.9368
6.4368 | -85.9373
6.4372 | -85.9376
6.4375 | -85.9375
6.4379 | -85.9382
6.4383 | -85.9404
6.4386 | -85.9416
6.439 | -85.9421
6.4393 | -85.9426
6.4397 | -85.9427
6.4401 | -85.9429
6.4404 | -85.9427
6.4408 | -85.943
6.4411 | -85.9438
6.4415 | -85.944
6.4419 | -85.9442
6.4422 | -85.945
6.4426 | -85.9454
6.443 | -85.9464
6.4433 | -85.9464
6.4437 | -85.9463
6.444 | -85.9469
6.4444 | -85.948
6.4448 | -85.9483
6.4451 | -85.9482
6.4455 | -85.948
6.4458 | -85.9476
6.4462 | -85.9477
6.4466 | -85.948
6.4469 | -85.9477
6.4473 | -85.9483
6.4477 | -85.9496
6.448 | -85.9494
6.4484 | -85.9491
6.4487 | -85.9493
6.4491 | -85.9491
6.4495 | -85.9479
6.4498 | -85.9473
6.4502 | -85.9473
6.4506 | -85.9475
6.4509 | -85.9464
6.4513 | -85.9464
6.4516 | -85.9461
6.452 | -85.9464
6.4524 | -85.9449
6.4527 | -85.9436
6.4531 | -85.9441
6.4534 | -85.9442
6.4538 | -85.9437
6.4542 | -85.9432
6.4545 | -85.9438
6.4549 | -85.945
6.4553 | -85.9469
6.4556 | -85.9509
6.456 | -85.953
6.4563 | -85.9534
6.4567 | -85.9536
6.4571 | -85.9534
6.4574 | -85.9534
6.4578 | -85.9527
6.4581 | -85.9524
6.4585 | -85.9525
6.4589 | -85.9529
6.4592 | -85.9527
6.4596 | -85.9525
6.46 | -85.9523
6.4603 | -85.9516
6.4607 | -85.9513
6.461 | -85.9514
6.4614 | -85.9512
6.4618 | -85.9518
6.4621 | -85.951
6.4625 | -85.9511
6.4629 | -85.9508
6.4632 | -85.9508
6.4636 | -85.951
6.4639 | -85.9509
6.4643 | -85.9506
6.4647 | -85.9509
6.465 | -85.9521
6.4654 | -85.9525
6.4657 | -85.9516
6.4661 | -85.9515
6.4665 | -85.9529
6.4668 | -85.9535
6.4672 | -85.9531
6.4676 | -85.953
6.4679 | -85.9529
6.4683 | -85.9527
6.4686 | -85.9522
6.469 | -85.9518
6.4694 | -85.9525
6.4697 | -85.953
6.4701 | -85.9536
6.4704 | -85.954
6.4708 | -85.9541
6.4712 | -85.9539
6.4715 | -85.9542
6.4719 | -85.9537
6.4723 | -85.956
6.4726 | -85.956
6.473 | -85.9562
6.4733 | -85.9557
6.4737 | -85.9554
6.4741 | -85.9552
6.4744 | -85.9554
6.4748 | -85.9547
6.4752 | -85.9543
6.4755 | -85.9539
6.4759 | -85.9531
6.4762 | -85.9517
6.4766 | -85.9507
6.477 | -85.9505
6.4773 | -85.9503
6.4777 | -85.95
6.478 | -85.9507
6.4784 | -85.9507
6.4788 | -85.9509
6.4791 | -85.9518
6.4795 | -85.9525
6.4799 | -85.9533
6.4802 | -85.9533
6.4806 | -85.9529
6.4809 | -85.9529
6.4813 | -85.9525
6.4817 | -85.9515
6.482 | -85.9507
6.4824 | -85.9506
6.4828 | -85.9503
6.4831 | -85.9504
6.4835 | -85.9512
6.4838 | -85.9515
6.4842 | -85.9506
6.4846 | -85.9501
6.4849 | -85.9496
6.4853 | -85.9483
6.4856 | -85.9476
6.486 | -85.9475
6.4864 | -85.947
6.4867 | -85.9465
6.4871 | -85.946
6.4875 | -85.9458
6.4878 | -85.9465
6.4882 | -85.9469
6.4885 | -85.9466
6.4889 | -85.9466
6.4893 | -85.9471
6.4896 | -85.9466
6.49 | -85.9462
6.4903 | -85.9444
6.4907 | -85.9425
6.4911 | -85.9428
6.4914 | -85.9433
6.4918 | -85.9431
6.4922 | -85.9425
6.4925 | -85.9426
6.4929 | -85.942
6.4932 | -85.942
6.4936 | -85.9405
6.494 | -85.9387
6.4943 | -85.938
6.4947 | -85.9376
6.4951 | -85.9372
6.4954 | -85.9376
6.4958 | -85.938
6.4961 | -85.9378
6.4965 | -85.9371
6.4969 | -85.9366
6.4972 | -85.9365
6.4976 | -85.9366
6.4979 | -85.9358
6.4983 | -85.9356
6.4987 | -85.9358
6.499 | -85.9347
6.4994 | -85.934
6.4998 | -85.9332
6.5001 | -85.9322
6.5005 | -85.9316
6.5008 | -85.9315
6.5012 | -85.9323
6.5016 | -85.9326
6.5019 | -85.9321
6.5023 | -85.9317
6.5027 | -85.9319
6.503 | -85.931
6.5034 | -85.9296
6.5037 | -85.9289
6.5041 | -85.9281
6.5045 | -85.9275
6.5048 | -85.9264
6.5052 | -85.9254
6.5055 | -85.9246
6.5059 | -85.924
6.5063 | -85.9236
6.5066 | -85.9229
6.507 | -85.922
6.5074 | -85.9215
6.5077 | -85.9208
6.5081 | -85.9201
6.5084 | -85.9196
6.5088 | -85.9191
6.5092 | -85.9186
6.5095 | -85.9179
6.5099 | -85.9172
6.5102 | -85.9157
6.5106 | -85.9146
6.511 | -85.9132
6.5113 | -85.912
6.5117 | -85.9111
6.5121 | -85.9101
6.5124 | -85.909
6.5128 | -85.9082
6.5131 | -85.9074
6.5135 | -85.9069
6.5139 | -85.9063
6.5142 | -85.9056
6.5146 | -85.9047
6.515 | -85.9044
6.5153 | -85.9046
6.5157 | -85.9042
6.516 | -85.9032
6.5164 | -85.9031
6.5168 | -85.9029
6.5171 | -85.9018
6.5175 | -85.9008
6.5178 | -85.9002
6.5182 | -85.9
6.5186 | -85.8996
6.5189 | -85.8996
6.5193 | -85.9003
6.5197 | -85.9014
6.52 | -85.9011
6.5204 | -85.9002
6.5207 | -85.8996
6.5211 | -85.8989
6.5215 | -85.8982
6.5218 | -85.8976
6.5222 | -85.8969
6.5225 | -85.8964
6.5229 | -85.8952
6.5233 | -85.8952
6.5236 | -85.8943
6.524 | -85.8943
6.5244 | -85.8932
6.5247 | -85.8916
6.5251 | -85.8909
6.5254 | -85.8903
6.5258 | -85.8895
6.5262 | -85.8887
6.5265 | -85.8879
6.5269 | -85.8869
6.5273 | -85.8857
6.5276 | -85.8848
6.528 | -85.884
6.5283 | -85.8828
6.5287 | -85.8825
6.5291 | -85.8813
6.5294 | -85.8809
6.5298 | -85.8803
6.5301 | -85.8799
6.5305 | -85.8795
6.5309 | -85.8791
6.5312 | -85.8792
6.5316 | -85.8788
6.532 | -85.8777
6.5323 | -85.8771
6.5327 | -85.8765
6.533 | -85.8772
6.5334 | -85.877
6.5338 | -85.8763
6.5341 | -85.8763
6.5345 | -85.8762
6.5348 | -85.8761
6.5352 | -85.8755
6.5356 | -85.8747
6.5359 | -85.8744
6.5363 | -85.8747
6.5367 | -85.8737
6.537 | -85.8731
6.5374 | -85.8725
6.5377 | -85.8713
6.5381 | -85.8709
6.5385 | -85.871
6.5388 | -85.8704
6.5392 | -85.8701
6.5396 | -85.8701
6.5399 | -85.8702
6.5403 | -85.8698
6.5406 | -85.8695
6.541 | -85.8698
6.5414 | -85.8695
6.5417 | -85.8686
6.5421 | -85.8683
6.5424 | -85.8679
6.5428 | -85.8686
6.5432 | -85.8694
6.5435 | -85.8691
6.5439 | -85.8685
6.5443 | -85.8677
6.5446 | -85.8665
6.545 | -85.8651
6.5453 | -85.865
6.5457 | -85.8645
6.5461 | -85.8642
6.5464 | -85.8632
6.5468 | -85.8625
6.5472 | -85.8614
6.5475 | -85.8598
6.5479 | -85.8594
6.5482 | -85.8588
6.5486 | -85.8582
6.549 | -85.8569
6.5493 | -85.8563
6.5497 | -85.8557
6.55 | -85.8544
6.5504 | -85.8527
6.5508 | -85.8514
6.5511 | -85.8507
6.5515 | -85.8498
6.5519 | -85.8486
6.5522 | -85.847
6.5526 | -85.8463
6.5529 | -85.846
6.5533 | -85.8452
6.5537 | -85.8443
6.554 | -85.8433
6.5544 | -85.8422
6.5548 | -85.8416
6.5551 | -85.841
6.5555 | -85.8397
6.5558 | -85.8397
6.5562 | -85.839
6.5566 | -85.8391
6.5569 | -85.8374
6.5573 | -85.8365
6.5576 | -85.8355
6.558 | -85.835
6.5584 | -85.8343
6.5587 | -85.834
6.5591 | -85.8336
6.5595 | -85.8329
6.5598 | -85.8325
6.5602 | -85.8322
6.5605 | -85.8321
6.5609 | -85.8317
6.5613 | -85.8318
6.5616 | -85.834
6.562 | -85.8346
6.5623 | -85.834
6.5627 | -85.833
6.5631 | -85.8321
6.5634 | -85.8319
6.5638 | -85.8318
6.5642 | -85.8314
6.5645 | -85.8305
6.5649 | -85.8309
6.5652 | -85.8307
6.5656 | -85.8304
6.566 | -85.8301
6.5663 | -85.8297
6.5667 | -85.8291
6.5671 | -85.8292
6.5674 | -85.8289
6.5678 | -85.8286
6.5681 | -85.8282
6.5685 | -85.8277
6.5689 | -85.8267
6.5692 | -85.8237
6.5696 | -85.8202
6.5699 | -85.8175
6.5703 | -85.8152
6.5707 | -85.8144
6.571 | -85.8138
6.5714 | -85.8132
6.5718 | -85.8126
6.5721 | -85.8121
6.5725 | -85.8122
6.5728 | -85.8117
6.5732 | -85.8099
6.5736 | -85.8086
6.5739 | -85.8071
6.5743 | -85.806
6.5746 | -85.8065
6.575 | -85.8068
6.5754 | -85.806
6.5757 | -85.805
6.5761 | -85.804
6.5765 | -85.802
6.5768 | -85.8029
6.5772 | -85.8027
6.5775 | -85.8023
6.5779 | -85.8013
6.5783 | -85.7993
6.5786 | -85.7976
6.579 | -85.7969
6.5794 | -85.7962
6.5797 | -85.7949
6.5801 | -85.7934
6.5804 | -85.7931
6.5808 | -85.7928
6.5812 | -85.7919
6.5815 | -85.7912
6.5819 | -85.7907
6.5822 | -85.7905
6.5826 | -85.7899
6.583 | -85.7884
6.5833 | -85.7875
6.5837 | -85.7871
6.5841 | -85.787
6.5844 | -85.7863
6.5848 | -85.7859
6.5851 | -85.7849
6.5855 | -85.7831
6.5859 | -85.782
6.5862 | -85.7814
6.5866 | -85.7804
6.5869 | -85.7795
6.5873 | -85.7788
6.5877 | -85.7782
6.588 | -85.7781
6.5884 | -85.7778
6.5888 | -85.7778
6.5891 | -85.7772
6.5895 | -85.7767
6.5898 | -85.7761
6.5902 | -85.7754
6.5906 | -85.7751
6.5909 | -85.7753
6.5913 | -85.7751
6.5917 | -85.7744
6.592 | -85.7739
6.5924 | -85.7733
6.5927 | -85.7727
6.5931 | -85.7714
6.5935 | -85.7706
6.5938 | -85.7717
6.5942 | -85.7726
6.5945 | -85.7727
6.5949 | -85.7727
6.5953 | -85.7728
6.5956 | -85.7727
6.596 | -85.7722
6.5964 | -85.7721
6.5967 | -85.7721
6.5971 | -85.7725
6.5974 | -85.772
6.5978 | -85.7721
6.5982 | -85.7738
6.5985 | -85.7747
6.5989 | -85.7743
6.5993 | -85.772
6.5996 | -85.7705
6.6 | -85.7701
6.6003 | -85.7703
6.6007 | -85.77
6.6011 | -85.7692
6.6014 | -85.768
6.6018 | -85.7667
6.6021 | -85.7661
6.6025 | -85.7666
6.6029 | -85.7671
6.6032 | -85.7669
6.6036 | -85.7667
6.604 | -85.7663
6.6043 | -85.7656
6.6047 | -85.7647
6.605 | -85.7647
6.6054 | -85.7641
6.6058 | -85.7631
6.6061 | -85.7635
6.6065 | -85.7631
6.6068 | -85.7627
6.6072 | -85.7626
6.6076 | -85.7622
6.6079 | -85.7621
6.6083 | -85.7616
6.6087 | -85.7611
6.609 | -85.7607
6.6094 | -85.7601
6.6097 | -85.7591
6.6101 | -85.7577
6.6105 | -85.7569
6.6108 | -85.7568
6.6112 | -85.7563
6.6116 | -85.7555
6.6119 | -85.7548
6.6123 | -85.7542
6.6126 | -85.7522
6.613 | -85.7513
6.6134 | -85.7511
6.6137 | -85.7507
6.6141 | -85.75
6.6144 | -85.7492
6.6148 | -85.7484
6.6152 | -85.7492
6.6155 | -85.7495
6.6159 | -85.7492
6.6163 | -85.7485
6.6166 | -85.7478
6.617 | -85.7468
6.6173 | -85.7459
6.6177 | -85.7454
6.6181 | -85.7445
6.6184 | -85.7442
6.6188 | -85.7443
6.6192 | -85.7445
6.6195 | -85.7441
6.6199 | -85.7437
6.6202 | -85.7436
6.6206 | -85.7429
6.621 | -85.7429
6.6213 | -85.7419
6.6217 | -85.7412
6.622 | -85.7407
6.6224 | -85.7406
6.6228 | -85.7413
6.6231 | -85.741
6.6235 | -85.7407
6.6239 | -85.74
6.6242 | -85.7396
6.6246 | -85.7396
6.6249 | -85.7394
6.6253 | -85.7389
6.6257 | -85.7386
6.626 | -85.7382
6.6264 | -85.7366
6.6267 | -85.735
6.6271 | -85.7342
6.6275 | -85.7335
6.6278 | -85.7322
6.6282 | -85.7312
6.6286 | -85.7313
6.6289 | -85.7313
6.6293 | -85.7315
6.6296 | -85.7312
6.63 | -85.7311
6.6304 | -85.7309
6.6307 | -85.7307
6.6311 | -85.7299
6.6315 | -85.7293
6.6318 | -85.729
6.6322 | -85.7288
6.6325 | -85.7286
6.6329 | -85.7276
6.6333 | -85.7266
6.6336 | -85.7264
6.634 | -85.7265
6.6343 | -85.7264
6.6347 | -85.7265
6.6351 | -85.7261
6.6354 | -85.7258
6.6358 | -85.7259
6.6362 | -85.7258
6.6365 | -85.7249
6.6369 | -85.7243
6.6372 | -85.724
6.6376 | -85.724
6.638 | -85.7243
6.6383 | -85.7238
6.6387 | -85.7234
6.639 | -85.7234
6.6394 | -85.7229
6.6398 | -85.7228
6.6401 | -85.722
6.6405 | -85.7211
6.6409 | -85.7198
6.6412 | -85.7191
6.6416 | -85.7191
6.6419 | -85.7192
6.6423 | -85.7191
6.6427 | -85.7191
6.643 | -85.7194
6.6434 | -85.7193
6.6438 | -85.7193
6.6441 | -85.7189
6.6445 | -85.7182
6.6448 | -85.7181
6.6452 | -85.7168
6.6456 | -85.7161
6.6459 | -85.716
6.6463 | -85.7158
6.6466 | -85.7155
6.647 | -85.7141
6.6474 | -85.714
6.6477 | -85.7149
6.6481 | -85.715
6.6485 | -85.7148
6.6488 | -85.7145
6.6492 | -85.7144
6.6495 | -85.7149
6.6499 | -85.7145
6.6503 | -85.7143
6.6506 | -85.7142
6.651 | -85.7143
6.6513 | -85.7142
6.6517 | -85.7141
6.6521 | -85.7139
6.6524 | -85.7139
6.6528 | -85.7141
6.6532 | -85.7145
6.6535 | -85.7148
6.6539 | -85.714
6.6542 | -85.7131
6.6546 | -85.7132
6.655 | -85.7141
6.6553 | -85.7142
6.6557 | -85.7137
6.6561 | -85.7139
6.6564 | -85.7141
6.6568 | -85.7142
6.6571 | -85.7142
6.6575 | -85.714
6.6579 | -85.7141
6.6582 | -85.7145
6.6586 | -85.7144
6.6589 | -85.715
6.6593 | -85.716
6.6597 | -85.7154
6.66 | -85.7156
6.6604 | -85.715
6.6608 | -85.7142
6.6611 | -85.7139
6.6615 | -85.714
6.6618 | -85.7139
6.6622 | -85.7137
6.6626 | -85.7136
6.6629 | -85.7135
6.6633 | -85.7146
6.6637 | -85.7161
6.664 | -85.7167
6.6644 | -85.7166
6.6647 | -85.7172
6.6651 | -85.7165
6.6655 | -85.716
6.6658 | -85.7162
6.6662 | -85.7168
6.6665 | -85.7167
6.6669 | -85.7171
6.6673 | -85.7166
6.6676 | -85.7157
6.668 | -85.7158
6.6684 | -85.7162
6.6687 | -85.716
6.6691 | -85.7149
6.6694 | -85.7149
6.6698 | -85.715
6.6702 | -85.7146
6.6705 | -85.7142
6.6709 | -85.714
6.6712 | -85.7138
6.6716 | -85.7137
6.672 | -85.7137
6.6723 | -85.7138
6.6727 | -85.7138
6.6731 | -85.7137
6.6734 | -85.714
6.6738 | -85.7134
6.6741 | -85.7127
6.6745 | -85.7126
6.6749 | -85.7123
6.6752 | -85.7119
6.6756 | -85.7118
6.676 | -85.7117
6.6763 | -85.7119
6.6767 | -85.7119
6.677 | -85.7118
6.6774 | -85.7107
6.6778 | -85.7102
6.6781 | -85.7102
6.6785 | -85.7105
6.6788 | -85.7112
6.6792 | -85.7119
6.6796 | -85.7121
6.6799 | -85.7118
6.6803 | -85.7113
6.6807 | -85.7109
6.681 | -85.7111
6.6814 | -85.7108
6.6817 | -85.7103
6.6821 | -85.7102
6.6825 | -85.7111
6.6828 | -85.7115
6.6832 | -85.7106
6.6836 | -85.7103
6.6839 | -85.7104
6.6843 | -85.7105
6.6846 | -85.7117
6.685 | -85.7133
6.6854 | -85.7147
6.6857 | -85.7157
6.6861 | -85.7161
6.6864 | -85.7159
6.6868 | -85.716
6.6872 | -85.7162
6.6875 | -85.7166
6.6879 | -85.717
6.6883 | -85.7171
6.6886 | -85.7171
6.689 | -85.7171
6.6893 | -85.7167
6.6897 | -85.7166
6.6901 | -85.7168
6.6904 | -85.7173
6.6908 | -85.7179
6.6911 | -85.7187
6.6915 | -85.7192
6.6919 | -85.7194
6.6922 | -85.719
6.6926 | -85.7184
6.693 | -85.7183
6.6933 | -85.7187
6.6937 | -85.7191
6.694 | -85.7192
6.6944 | -85.72
6.6948 | -85.7202
6.6951 | -85.7205
6.6955 | -85.7214
6.6959 | -85.7214
6.6962 | -85.7204
6.6966 | -85.7185
6.6969 | -85.717
6.6973 | -85.7165
6.6977 | -85.7164
6.698 | -85.7167
6.6984 | -85.7169
6.6987 | -85.7168
6.6991 | -85.7167
6.6995 | -85.7168
6.6998 | -85.7171
6.7002 | -85.7169
6.7006 | -85.7175
6.7009 | -85.7175
6.7013 | -85.7181
6.7016 | -85.7184
6.702 | -85.7188
6.7024 | -85.7193
6.7027 | -85.7194
6.7031 | -85.7196
6.7035 | -85.7199
6.7038 | -85.7194
6.7042 | -85.7199
6.7045 | -85.7204
6.7049 | -85.7204
6.7053 | -85.7202
6.7056 | -85.7196
6.706 | -85.7194
6.7063 | -85.7198
6.7067 | -85.7198
6.7071 | -85.7191
6.7074 | -85.7187
6.7078 | -85.7181
6.7082 | -85.7182
6.7085 | -85.7183
6.7089 | -85.7184
6.7092 | -85.7189
6.7096 | -85.7194
6.71 | -85.7189
6.7103 | -85.7191
6.7107 | -85.7195
6.711 | -85.7192
6.7114 | -85.7195
6.7118 | -85.7201
6.7121 | -85.7207
6.7125 | -85.7207
6.7129 | -85.7209
6.7132 | -85.7204
6.7136 | -85.7199
6.7139 | -85.7202
6.7143 | -85.7204
6.7147 | -85.7202
6.715 | -85.7204
6.7154 | -85.7212
6.7158 | -85.7209
6.7161 | -85.7201
6.7165 | -85.7184
6.7168 | -85.7179
6.7172 | -85.7183
6.7176 | -85.7183
6.7179 | -85.7186
6.7183 | -85.7192
6.7186 | -85.7199
6.719 | -85.72
6.7194 | -85.7195
6.7197 | -85.7195
6.7201 | -85.7196
6.7205 | -85.7198
6.7208 | -85.72
6.7212 | -85.7201
6.7215 | -85.7208
6.7219 | -85.7211
6.7223 | -85.7218
6.7226 | -85.7222
6.723 | -85.723
6.7233 | -85.7234
6.7237 | -85.724
6.7241 | -85.7245
6.7244 | -85.7248
6.7248 | -85.725
6.7252 | -85.725
6.7255 | -85.7253
6.7259 | -85.7258
6.7262 | -85.7253
6.7266 | -85.725
6.727 | -85.7247
6.7273 | -85.7252
6.7277 | -85.7255
6.7281 | -85.7259
6.7284 | -85.7254
6.7288 | -85.7242
6.7291 | -85.7237
6.7295 | -85.7234
6.7299 | -85.7235
6.7302 | -85.7234
6.7306 | -85.7234
6.7309 | -85.7236
6.7313 | -85.724
6.7317 | -85.724
6.732 | -85.7239
6.7324 | -85.7243
6.7328 | -85.725
6.7331 | -85.7254
6.7335 | -85.7259
6.7338 | -85.7261
6.7342 | -85.7265
6.7346 | -85.7266
6.7349 | -85.7274
6.7353 | -85.7288
6.7356 | -85.7296
6.736 | -85.7306
6.7364 | -85.7312
6.7367 | -85.7313
6.7371 | -85.7314
6.7375 | -85.7309
6.7378 | -85.7312
6.7382 | -85.7314
6.7385 | -85.7315
6.7389 | -85.7314
6.7393 | -85.7316
6.7396 | -85.7317
6.74 | -85.7312
6.7404 | -85.7317
6.7407 | -85.732
6.7411 | -85.7322
6.7414 | -85.7327
6.7418 | -85.733
6.7422 | -85.7333
6.7425 | -85.7333
6.7429 | -85.7333
6.7432 | -85.7342
6.7436 | -85.7335
6.744 | -85.7324
6.7443 | -85.7326
6.7447 | -85.7325
6.7451 | -85.7322
6.7454 | -85.7332
6.7458 | -85.7337
6.7461 | -85.7347
6.7465 | -85.7352
6.7469 | -85.7353
6.7472 | -85.7354
6.7476 | -85.7351
6.748 | -85.7344
6.7483 | -85.7338
6.7487 | -85.7337
6.749 | -85.7339
6.7494 | -85.7337
6.7498 | -85.7336
6.7501 | -85.7333
6.7505 | -85.7326
6.7508 | -85.7329
6.7512 | -85.7334
6.7516 | -85.7334
6.7519 | -85.7337
6.7523 | -85.7336
6.7527 | -85.7339
6.753 | -85.7342
6.7534 | -85.7342
6.7537 | -85.7342
6.7541 | -85.7338
6.7545 | -85.7338
6.7548 | -85.7334
6.7552 | -85.7336
6.7555 | -85.7339
6.7559 | -85.7342
6.7563 | -85.7348
6.7566 | -85.7344
6.757 | -85.7347
6.7574 | -85.7352
6.7577 | -85.7353
6.7581 | -85.7356
6.7584 | -85.7355
6.7588 | -85.7354
6.7592 | -85.7355
6.7595 | -85.736
6.7599 | -85.7362
6.7603 | -85.7362
6.7606 | -85.7368
6.761 | -85.7378
6.7613 | -85.738
6.7617 | -85.7379
6.7621 | -85.7376
6.7624 | -85.7377
6.7628 | -85.7377
6.7631 | -85.7382
6.7635 | -85.7396
6.7639 | -85.7406
6.7642 | -85.7393
6.7646 | -85.739
6.765 | -85.7389
6.7653 | -85.7393
6.7657 | -85.7385
6.766 | -85.7384
6.7664 | -85.739
6.7668 | -85.7399
6.7671 | -85.7406
6.7675 | -85.7401
6.7679 | -85.7405
6.7682 | -85.7412
6.7686 | -85.7416
6.7689 | -85.7424
6.7693 | -85.7427
6.7697 | -85.743
6.77 | -85.7433
6.7704 | -85.7434
6.7707 | -85.7435
6.7711 | -85.7433
6.7715 | -85.7436
6.7718 | -85.7441
6.7722 | -85.7443
6.7726 | -85.7442
6.7729 | -85.744
6.7733 | -85.744
6.7736 | -85.7447
6.774 | -85.7455
6.7744 | -85.7461
6.7747 | -85.7469
6.7751 | -85.7472
6.7754 | -85.7483
6.7758 | -85.7491
6.7762 | -85.7496
6.7765 | -85.7497
6.7769 | -85.7499
6.7773 | -85.7502
6.7776 | -85.7505
6.778 | -85.7507
6.7783 | -85.7513
6.7787 | -85.7522
6.7791 | -85.7527
6.7794 | -85.7538
6.7798 | -85.7543
6.7802 | -85.7545
6.7805 | -85.7541
6.7809 | -85.7536
6.7812 | -85.7542
6.7816 | -85.7541
6.782 | -85.7541
6.7823 | -85.7548
6.7827 | -85.754
6.783 | -85.7539
6.7834 | -85.7543
6.7838 | -85.7553
6.7841 | -85.7558
6.7845 | -85.7555
6.7849 | -85.7561
6.7852 | -85.7566
6.7856 | -85.7572
6.7859 | -85.7569
6.7863 | -85.7568
6.7867 | -85.7581
6.787 | -85.7589
6.7874 | -85.7585
6.7877 | -85.7583
6.7881 | -85.7585
6.7885 | -85.7593
6.7888 | -85.7594
6.7892 | -85.7597
6.7896 | -85.7599
6.7899 | -85.7599
6.7903 | -85.76
6.7906 | -85.7603
6.791 | -85.7611
6.7914 | -85.7614
6.7917 | -85.7618
6.7921 | -85.7624
6.7925 | -85.7625
6.7928 | -85.7625
6.7932 | -85.7629
6.7935 | -85.7632
6.7939 | -85.7654
6.7943 | -85.7664
6.7946 | -85.7659
6.795 | -85.7658
6.7953 | -85.7669
6.7957 | -85.7678
6.7961 | -85.7709
6.7964 | -85.773
6.7968 | -85.7735
6.7972 | -85.7739
6.7975 | -85.7743
6.7979 | -85.7746
6.7982 | -85.7754
6.7986 | -85.7768
6.799 | -85.7771
6.7993 | -85.7777
6.7997 | -85.7779
6.8001 | -85.7779
6.8004 | -85.7776
6.8008 | -85.7779
6.8011 | -85.7784
6.8015 | -85.7789
6.8019 | -85.7796
6.8022 | -85.7802
6.8026 | -85.7803
6.8029 | -85.7807
6.8033 | -85.781
6.8037 | -85.7822
6.804 | -85.783
6.8044 | -85.7835
6.8048 | -85.7844
6.8051 | -85.7851
6.8055 | -85.7857
6.8058 | -85.7866
6.8062 | -85.7881
6.8066 | -85.788
6.8069 | -85.7879
6.8073 | -85.7882
6.8076 | -85.7884
6.808 | -85.7883
6.8084 | -85.7882
6.8087 | -85.7884
6.8091 | -85.7889
6.8095 | -85.7895
6.8098 | -85.79
6.8102 | -85.7905
6.8105 | -85.7917
6.8109 | -85.7926
6.8113 | -85.793
6.8116 | -85.7928
6.812 | -85.7931
6.8124 | -85.7934
6.8127 | -85.7929
6.8131 | -85.7931
6.8134 | -85.7935
6.8138 | -85.794
6.8142 | -85.7942
6.8145 | -85.796
6.8149 | -85.797
6.8152 | -85.7976
6.8156 | -85.7985
6.816 | -85.7989
6.8163 | -85.7998
6.8167 | -85.8004
6.8171 | -85.8006
6.8174 | -85.8009
6.8178 | -85.8013
6.8181 | -85.8021
6.8185 | -85.8028
6.8189 | -85.8038
6.8192 | -85.8046
6.8196 | -85.8052
6.82 | -85.8052
6.8203 | -85.8064
6.8207 | -85.8074
6.821 | -85.8076
6.8214 | -85.8081
6.8218 | -85.8086
6.8221 | -85.8093
6.8225 | -85.8101
6.8228 | -85.8109
6.8232 | -85.8115
6.8236 | -85.8116
6.8239 | -85.8123
6.8243 | -85.8133
6.8247 | -85.8148
6.825 | -85.8159
6.8254 | -85.8164
6.8257 | -85.8164
6.8261 | -85.8161
6.8265 | -85.8162
6.8268 | -85.8175
6.8272 | -85.818
6.8275 | -85.8185
6.8279 | -85.8192
6.8283 | -85.8203
6.8286 | -85.8204
6.829 | -85.821
6.8294 | -85.8215
6.8297 | -85.8212
6.8301 | -85.8211
6.8304 | -85.8217
6.8308 | -85.8226
6.8312 | -85.8235
6.8315 | -85.8239
6.8319 | -85.824
6.8323 | -85.8243
6.8326 | -85.8248
6.833 | -85.8252
6.8333 | -85.8254
6.8337 | -85.8261
6.8341 | -85.8268
6.8344 | -85.8275
6.8348 | -85.8281
6.8351 | -85.8277
6.8355 | -85.8277
6.8359 | -85.8281
6.8362 | -85.828
6.8366 | -85.8281
6.837 | -85.8287
6.8373 | -85.8301
6.8377 | -85.8303
6.838 | -85.8308
6.8384 | -85.8315
6.8388 | -85.8319
6.8391 | -85.8325
6.8395 | -85.8332
6.8398 | -85.8338
6.8402 | -85.8343
6.8406 | -85.835
6.8409 | -85.8359
6.8413 | -85.8367
6.8417 | -85.8374
6.842 | -85.8388
6.8424 | -85.84
6.8427 | -85.8404
6.8431 | -85.8414
6.8435 | -85.8432
6.8438 | -85.844
6.8442 | -85.8437
6.8446 | -85.844
6.8449 | -85.8446
6.8453 | -85.845
6.8456 | -85.8457
6.846 | -85.8462
6.8464 | -85.8461
6.8467 | -85.8463
6.8471 | -85.847
6.8474 | -85.8478
6.8478 | -85.8493
6.8482 | -85.8505
6.8485 | -85.8521
6.8489 | -85.8531
6.8493 | -85.8542
6.8496 | -85.8554
6.85 | -85.8561
6.8503 | -85.8568
6.8507 | -85.8574
6.8511 | -85.8583
6.8514 | -85.8595
6.8518 | -85.8608
6.8521 | -85.8616
6.8525 | -85.8628
6.8529 | -85.8638
6.8532 | -85.8655
6.8536 | -85.8671
6.854 | -85.8686
6.8543 | -85.8701
6.8547 | -85.8717
6.855 | -85.8722
6.8554 | -85.8728
6.8558 | -85.8734
6.8561 | -85.8739
6.8565 | -85.8744
6.8569 | -85.8748
6.8572 | -85.8756
6.8576 | -85.8764
6.8579 | -85.8771
6.8583 | -85.8767
6.8587 | -85.8767
6.859 | -85.8775
6.8594 | -85.8779
6.8597 | -85.8782
6.8601 | -85.8789
6.8605 | -85.8798
6.8608 | -85.8813
6.8612 | -85.8832
6.8616 | -85.8859
6.8619 | -85.8884
6.8623 | -85.8895
6.8626 | -85.8897
6.863 | -85.8904
6.8634 | -85.8918
6.8637 | -85.893
6.8641 | -85.8937
6.8645 | -85.896
6.8648 | -85.8988
6.8652 | -85.9002
6.8655 | -85.9009
6.8659 | -85.9018
6.8663 | -85.9029
6.8666 | -85.9046
6.867 | -85.9061
6.8673 | -85.9068
6.8677 | -85.9076
6.8681 | -85.9082
6.8684 | -85.9093
6.8688 | -85.9103
6.8692 | -85.9121
6.8695 | -85.9139
6.8699 | -85.9158
6.8702 | -85.9174
6.8706 | -85.9185
6.871 | -85.9195
6.8713 | -85.9206
6.8717 | -85.922
6.872 | -85.9238
6.8724 | -85.925
6.8728 | -85.9257
6.8731 | -85.9264
6.8735 | -85.9271
6.8739 | -85.9274
6.8742 | -85.9278
6.8746 | -85.9279
6.8749 | -85.9278
6.8753 | -85.9286
6.8757 | -85.9294
6.876 | -85.9308
6.8764 | -85.9318
6.8768 | -85.9327
6.8771 | -85.9342
6.8775 | -85.9355
6.8778 | -85.9366
6.8782 | -85.9375
6.8786 | -85.9385
6.8789 | -85.9402
6.8793 | -85.9413
6.8796 | -85.9426
6.88 | -85.944
6.8804 | -85.945
6.8807 | -85.9468
6.8811 | -85.9485
6.8815 | -85.9493
6.8818 | -85.9501
6.8822 | -85.9511
6.8825 | -85.9519
6.8829 | -85.9538
6.8833 | -85.9548
6.8836 | -85.956
6.884 | -85.9573
6.8844 | -85.9582
6.8847 | -85.959
6.8851 | -85.96
6.8854 | -85.9623
6.8858 | -85.9638
6.8862 | -85.9643
6.8865 | -85.9655
6.8869 | -85.9672
6.8872 | -85.9694
6.8876 | -85.971
6.888 | -85.9721
6.8883 | -85.9723
6.8887 | -85.973
6.8891 | -85.9739
6.8894 | -85.9748
6.8898 | -85.9764
6.8901 | -85.9778
6.8905 | -85.979
6.8909 | -85.9814
6.8912 | -85.9829
6.8916 | -85.9836
6.8919 | -85.9853
6.8923 | -85.9867
6.8927 | -85.988
6.893 | -85.989
6.8934 | -85.9905
6.8938 | -85.9922
6.8941 | -85.9933
6.8945 | -85.9945
6.8948 | -85.9959
6.8952 | -85.9974
6.8956 | -85.9985
6.8959 | -85.9998
6.8963 | -86.0011
6.8967 | -86.0021
6.897 | -86.0035
6.8974 | -86.0052
6.8977 | -86.0071
6.8981 | -86.0083
6.8985 | -86.0097
6.8988 | -86.0118
6.8992 | -86.0129
6.8995 | -86.0133
6.8999 | -86.0145
6.9003 | -86.0156
6.9006 | -86.017
6.901 | -86.0183
6.9014 | -86.0196
6.9017 | -86.0207
6.9021 | -86.0226
6.9024 | -86.024
6.9028 | -86.0255
6.9032 | -86.0268
6.9035 | -86.0275
6.9039 | -86.0285
6.9042 | -86.0298
6.9046 | -86.031
6.905 | -86.0326
6.9053 | -86.034
6.9057 | -86.0352
6.9061 | -86.0365
6.9064 | -86.0379
6.9068 | -86.0395
6.9071 | -86.0413
6.9075 | -86.0425
6.9079 | -86.0428
6.9082 | -86.0435
6.9086 | -86.0447
6.909 | -86.0456
6.9093 | -86.0468
6.9097 | -86.0481
6.91 | -86.0495
6.9104 | -86.0515
6.9108 | -86.0524
6.9111 | -86.053
6.9115 | -86.0537
6.9118 | -86.0548
6.9122 | -86.0554
6.9126 | -86.056
6.9129 | -86.0574
6.9133 | -86.0595
6.9137 | -86.0612
6.914 | -86.0624
6.9144 | -86.0644
6.9147 | -86.0664
6.9151 | -86.0682
6.9155 | -86.0701
6.9158 | -86.0718
6.9162 | -86.0739
6.9166 | -86.077
6.9169 | -86.0796
6.9173 | -86.0811
6.9176 | -86.0826
6.918 | -86.0845
6.9184 | -86.0865
6.9187 | -86.088
6.9191 | -86.0893
6.9194 | -86.0904
6.9198 | -86.0916
6.9202 | -86.0929
6.9205 | -86.0942
6.9209 | -86.0964
6.9213 | -86.0981
6.9216 | -86.0996
6.922 | -86.1012
6.9223 | -86.1033
6.9227 | -86.1052
6.9231 | -86.1066
6.9234 | -86.1082
6.9238 | -86.1102
6.9241 | -86.1121
6.9245 | -86.1138
6.9249 | -86.1162
6.9252 | -86.118
6.9256 | -86.119
6.926 | -86.12
6.9263 | -86.1213
6.9267 | -86.1227
6.927 | -86.1246
6.9274 | -86.1261
6.9278 | -86.1276
6.9281 | -86.1295
6.9285 | -86.1309
6.9289 | -86.1318
6.9292 | -86.1326
6.9296 | -86.1342
6.9299 | -86.1363
6.9303 | -86.1375
6.9307 | -86.139
6.931 | -86.1402
6.9314 | -86.1425
6.9317 | -86.1436
6.9321 | -86.1442
6.9325 | -86.1457
6.9328 | -86.1473
6.9332 | -86.1487
6.9336 | -86.1504
6.9339 | -86.1523
6.9343 | -86.1542
6.9346 | -86.156
6.935 | -86.1578
6.9354 | -86.1599
6.9357 | -86.1622
6.9361 | -86.1639
6.9364 | -86.1654
6.9368 | -86.1673
6.9372 | -86.1686
6.9375 | -86.1697
6.9379 | -86.1707
6.9383 | -86.1726
6.9386 | -86.1739
6.939 | -86.1754
6.9393 | -86.1759
6.9397 | -86.1763
6.9401 | -86.1773
6.9404 | -86.1785
6.9408 | -86.1783
6.9412 | -86.1791
6.9415 | -86.1811
6.9419 | -86.1825
6.9422 | -86.185
6.9426 | -86.1871
6.943 | -86.1888
6.9433 | -86.1902
6.9437 | -86.1916
6.944 | -86.1933
6.9444 | -86.1953
6.9448 | -86.1968
6.9451 | -86.1985
6.9455 | -86.2001
6.9459 | -86.2011
6.9462 | -86.2027
6.9466 | -86.204
6.9469 | -86.2053
6.9473 | -86.2062
6.9477 | -86.2065
6.948 | -86.207
6.9484 | -86.2079
6.9488 | -86.209
6.9491 | -86.2103
6.9495 | -86.2119
6.9498 | -86.2137
6.9502 | -86.2153
6.9506 | -86.2177
6.9509 | -86.2191
6.9513 | -86.219
6.9516 | -86.2207
6.952 | -86.2221
6.9524 | -86.2227
6.9527 | -86.2237
6.9531 | -86.2257
6.9535 | -86.2272
6.9538 | -86.229
6.9542 | -86.231
6.9545 | -86.2325
6.9549 | -86.2339
6.9553 | -86.2358
6.9556 | -86.2379
6.956 | -86.2401
6.9563 | -86.2421
6.9567 | -86.2435
6.9571 | -86.2464
6.9574 | -86.2498
6.9578 | -86.2519
6.9582 | -86.2536
6.9585 | -86.2555
6.9589 | -86.258
6.9592 | -86.2601
6.9596 | -86.2621
6.96 | -86.2636
6.9603 | -86.2654
6.9607 | -86.2679
6.9611 | -86.271
6.9614 | -86.273
6.9618 | -86.2745
6.9621 | -86.2758
6.9625 | -86.277
6.9629 | -86.2783
6.9632 | -86.2799
6.9636 | -86.2815
6.9639 | -86.2835
6.9643 | -86.285
6.9647 | -86.2867
6.965 | -86.2881
6.9654 | -86.29
6.9658 | -86.2918
6.9661 | -86.2937
6.9665 | -86.2961
6.9668 | -86.2982
6.9672 | -86.3001
6.9676 | -86.3017
6.9679 | -86.3032
6.9683 | -86.3048
6.9687 | -86.306
6.969 | -86.3073
6.9694 | -86.3086
6.9697 | -86.31
6.9701 | -86.3117
6.9705 | -86.3138
6.9708 | -86.3158
6.9712 | -86.318
6.9715 | -86.3191
6.9719 | -86.3205
6.9723 | -86.3221
6.9726 | -86.324
6.973 | -86.3261
6.9734 | -86.3281
6.9737 | -86.3298
6.9741 | -86.3313
6.9744 | -86.3329
6.9748 | -86.3347
6.9752 | -86.3371
6.9755 | -86.3392
6.9759 | -86.3407
6.9762 | -86.3422
6.9766 | -86.3438
6.977 | -86.3452
6.9773 | -86.3466
6.9777 | -86.3482
6.9781 | -86.3498
6.9784 | -86.3523
6.9788 | -86.3545
6.9791 | -86.357
6.9795 | -86.3599
6.9799 | -86.3621
6.9802 | -86.364
6.9806 | -86.3661
6.981 | -86.3677
6.9813 | -86.3688
6.9817 | -86.3699
6.982 | -86.3709
6.9824 | -86.372
6.9828 | -86.3738
6.9831 | -86.3757
6.9835 | -86.3759
6.9838 | -86.3767
6.9842 | -86.378
6.9846 | -86.3798
6.9849 | -86.382
6.9853 | -86.3833
6.9857 | -86.3849
6.986 | -86.3871
6.9864 | -86.3889
6.9867 | -86.3904
6.9871 | -86.3918
6.9875 | -86.393
6.9878 | -86.3948
6.9882 | -86.3964
6.9885 | -86.3978
6.9889 | -86.3994
6.9893 | -86.4013
6.9896 | -86.4029
6.99 | -86.4049
6.9904 | -86.4067
6.9907 | -86.4078
6.9911 | -86.4095
6.9914 | -86.4111
6.9918 | -86.4141
6.9922 | -86.416
6.9925 | -86.4174
6.9929 | -86.4183
6.9933 | -86.4192
6.9936 | -86.4207
6.994 | -86.4222
6.9943 | -86.425
6.9947 | -86.427
6.9951 | -86.4286
6.9954 | -86.4303
6.9958 | -86.4318
6.9961 | -86.4342
6.9965 | -86.436
6.9969 | -86.4377
6.9972 | -86.4396
6.9976 | -86.4414
6.998 | -86.4434
6.9983 | -86.4451
6.9987 | -86.447
6.999 | -86.448
6.9994 | -86.4492
6.9998 | -86.4508
7.0001 | -86.4525
7.0005 | -86.4534
7.0008 | -86.4546
7.0012 | -86.4565
7.0016 | -86.4586
7.0019 | -86.4602
7.0023 | -86.462
7.0027 | -86.4635
7.003 | -86.4652
7.0034 | -86.4665
7.0037 | -86.4677
7.0041 | -86.4695
7.0045 | -86.4713
7.0048 | -86.4729
7.0052 | -86.4745
7.0056 | -86.4762
7.0059 | -86.4777
7.0063 | -86.479
7.0066 | -86.4805
7.007 | -86.4823
7.0074 | -86.484
7.0077 | -86.4857
7.0081 | -86.4878
7.0084 | -86.4897
7.0088 | -86.4916
7.0092 | -86.4935
7.0095 | -86.4951
7.0099 | -86.4966
7.0103 | -86.4981
7.0106 | -86.4999
7.011 | -86.5015
7.0113 | -86.5032
7.0117 | -86.5053
7.0121 | -86.5075
7.0124 | -86.5091
7.0128 | -86.511
7.0132 | -86.5124
7.0135 | -86.513
7.0139 | -86.514
7.0142 | -86.5152
7.0146 | -86.5168
7.015 | -86.518
7.0153 | -86.5205
7.0157 | -86.5228
7.016 | -86.5247
7.0164 | -86.5262
7.0168 | -86.5279
7.0171 | -86.5299
7.0175 | -86.5313
7.0179 | -86.5321
7.0182 | -86.5342
7.0186 | -86.5365
7.0189 | -86.5374
7.0193 | -86.539
7.0197 | -86.5406
7.02 | -86.5427
7.0204 | -86.5443
7.0207 | -86.5464
7.0211 | -86.5483
7.0215 | -86.5501
7.0218 | -86.5524
7.0222 | -86.554
7.0226 | -86.5554
7.0229 | -86.5571
7.0233 | -86.5598
7.0236 | -86.5608
7.024 | -86.561
7.0244 | -86.5621
7.0247 | -86.5637
7.0251 | -86.5654
7.0255 | -86.5678
7.0258 | -86.5696
7.0262 | -86.5711
7.0265 | -86.5722
7.0269 | -86.5741
7.0273 | -86.576
7.0276 | -86.5772
7.028 | -86.5786
7.0283 | -86.581
7.0287 | -86.5841
7.0291 | -86.5864
7.0294 | -86.5884
7.0298 | -86.59
7.0302 | -86.5915
7.0305 | -86.5931
7.0309 | -86.5948
7.0312 | -86.5961
7.0316 | -86.5975
7.032 | -86.5989
7.0323 | -86.6
7.0327 | -86.6015
7.0331 | -86.6025
7.0334 | -86.6038
7.0338 | -86.6052
7.0341 | -86.6065
7.0345 | -86.6085
7.0349 | -86.6106
7.0352 | -86.6123
7.0356 | -86.6139
7.0359 | -86.6158
7.0363 | -86.6178
7.0367 | -86.6197
7.037 | -86.6212
7.0374 | -86.6227
7.0378 | -86.6245
7.0381 | -86.6263
7.0385 | -86.6278
7.0388 | -86.6295
7.0392 | -86.6304
7.0396 | -86.6311
7.0399 | -86.6325
7.0403 | -86.6341
7.0406 | -86.6357
7.041 | -86.6374
7.0414 | -86.6398
7.0417 | -86.6441
7.0421 | -86.649
7.0425 | -86.6522
7.0428 | -86.6551
7.0432 | -86.6571
7.0435 | -86.658
7.0439 | -86.6589
7.0443 | -86.6601
7.0446 | -86.6618
7.045 | -86.6636
7.0454 | -86.6657
7.0457 | -86.6686
7.0461 | -86.6698
7.0464 | -86.6727
7.0468 | -86.6754
7.0472 | -86.6782
7.0475 | -86.6799
7.0479 | -86.6817
7.0482 | -86.6837
7.0486 | -86.6855
7.049 | -86.6879
7.0493 | -86.6905
7.0497 | -86.6928
7.0501 | -86.6936
7.0504 | -86.6941
7.0508 | -86.6951
7.0511 | -86.6966
7.0515 | -86.6991
7.0519 | -86.7015
7.0522 | -86.7028
7.0526 | -86.704
7.053 | -86.7059
7.0533 | -86.7081
7.0537 | -86.7095
7.054 | -86.7111
7.0544 | -86.7127
7.0548 | -86.7138
7.0551 | -86.715
7.0555 | -86.7153
7.0558 | -86.7161
7.0562 | -86.7183
7.0566 | -86.7197
7.0569 | -86.7213
7.0573 | -86.7225
7.0577 | -86.7236
7.058 | -86.7247
7.0584 | -86.7257
7.0587 | -86.7275
7.0591 | -86.73
7.0595 | -86.7314
7.0598 | -86.7322
7.0602 | -86.7336
7.0605 | -86.736
7.0609 | -86.7376
7.0613 | -86.739
7.0616 | -86.7401
7.062 | -86.7409
7.0624 | -86.7419
7.0627 | -86.7426
7.0631 | -86.7434
7.0634 | -86.7444
7.0638 | -86.7454
7.0642 | -86.7469
7.0645 | -86.7479
7.0649 | -86.749
7.0653 | -86.7501
7.0656 | -86.7517
7.066 | -86.7531
7.0663 | -86.7544
7.0667 | -86.7561
7.0671 | -86.7574
7.0674 | -86.76
7.0678 | -86.7616
7.0681 | -86.7629
7.0685 | -86.7643
7.0689 | -86.7654
7.0692 | -86.7665
7.0696 | -86.7678
7.07 | -86.7691
7.0703 | -86.7706
7.0707 | -86.7723
7.071 | -86.7736
7.0714 | -86.7742
7.0718 | -86.7747
7.0721 | -86.7758
7.0725 | -86.7773
7.0728 | -86.7778
7.0732 | -86.7784
7.0736 | -86.7801
7.0739 | -86.7837
7.0743 | -86.7863
7.0747 | -86.7877
7.075 | -86.7892
7.0754 | -86.7906
7.0757 | -86.7924
7.0761 | -86.7945
7.0765 | -86.7968
7.0768 | -86.7986
7.0772 | -86.8
7.0776 | -86.8013
7.0779 | -86.8025
7.0783 | -86.8035
7.0786 | -86.8046
7.079 | -86.8059
7.0794 | -86.8072
7.0797 | -86.809
7.0801 | -86.8101
7.0804 | -86.8106
7.0808 | -86.8117
7.0812 | -86.8115
7.0815 | -86.8123
7.0819 | -86.8135
7.0823 | -86.8145
7.0826 | -86.8158
7.083 | -86.8178
7.0833 | -86.8189
7.0837 | -86.8202
7.0841 | -86.8214
7.0844 | -86.8223
7.0848 | -86.8237
7.0851 | -86.8256
7.0855 | -86.8279
7.0859 | -86.83
7.0862 | -86.8308
7.0866 | -86.8318
7.087 | -86.8334
7.0873 | -86.835
7.0877 | -86.8359
7.088 | -86.8372
7.0884 | -86.8403
7.0888 | -86.8417
7.0891 | -86.8423
7.0895 | -86.8433
7.0899 | -86.8445
7.0902 | -86.8461
7.0906 | -86.8478
7.0909 | -86.8489
7.0913 | -86.8501
7.0917 | -86.851
7.092 | -86.8527
7.0924 | -86.8542
7.0927 | -86.8558
7.0931 | -86.857
7.0935 | -86.8582
7.0938 | -86.8605
7.0942 | -86.8617
7.0946 | -86.8626
7.0949 | -86.8641
7.0953 | -86.8656
7.0956 | -86.8669
7.096 | -86.8673
7.0964 | -86.8683
7.0967 | -86.8697
7.0971 | -86.8706
7.0975 | -86.8724
7.0978 | -86.8736
7.0982 | -86.8739
7.0985 | -86.8748
7.0989 | -86.876
7.0993 | -86.8772
7.0996 | -86.8784
7.1 | -86.88
7.1003 | -86.8814
7.1007 | -86.8817
7.1011 | -86.8813
7.1014 | -86.8814
7.1018 | -86.8821
7.1022 | -86.8834
7.1025 | -86.8841
7.1029 | -86.8865
7.1032 | -86.8891
7.1036 | -86.8906
7.104 | -86.8918
7.1043 | -86.8936
7.1047 | -86.8947
7.105 | -86.8955
7.1054 | -86.897
7.1058 | -86.8987
7.1061 | -86.8995
7.1065 | -86.8997
7.1069 | -86.9002
7.1072 | -86.9012
7.1076 | -86.9021
7.1079 | -86.9041
7.1083 | -86.906
7.1087 | -86.907
7.109 | -86.9079
7.1094 | -86.9089
7.1098 | -86.91
7.1101 | -86.9119
7.1105 | -86.9139
7.1108 | -86.9148
7.1112 | -86.9155
7.1116 | -86.9161
7.1119 | -86.9166
7.1123 | -86.9178
7.1126 | -86.9186
7.113 | -86.9195
7.1134 | -86.9203
7.1137 | -86.922
7.1141 | -86.9236
7.1145 | -86.9248
7.1148 | -86.9258
7.1152 | -86.9267
7.1155 | -86.9277
7.1159 | -86.9286
7.1163 | -86.9296
7.1166 | -86.9312
7.117 | -86.9322
7.1174 | -86.9325
7.1177 | -86.934
7.1181 | -86.9357
7.1184 | -86.9376
7.1188 | -86.9389
7.1192 | -86.9391
7.1195 | -86.94
7.1199 | -86.9408
7.1202 | -86.9417
7.1206 | -86.943
7.121 | -86.9441
7.1213 | -86.9448
7.1217 | -86.9453
7.1221 | -86.9464
7.1224 | -86.9484
7.1228 | -86.9503
7.1231 | -86.9528
7.1235 | -86.9544
7.1239 | -86.9554
7.1242 | -86.9559
7.1246 | -86.9567
7.1249 | -86.9589
7.1253 | -86.96
7.1257 | -86.9604
7.126 | -86.9609
7.1264 | -86.9615
7.1268 | -86.9619
7.1271 | -86.963
7.1275 | -86.9638
7.1278 | -86.9647
7.1282 | -86.966
7.1286 | -86.9664
7.1289 | -86.9673
7.1293 | -86.9678
7.1297 | -86.968
7.13 | -86.9676
7.1304 | -86.9679
7.1307 | -86.9688
7.1311 | -86.969
7.1315 | -86.9695
7.1318 | -86.9703
7.1322 | -86.9716
7.1325 | -86.9737
7.1329 | -86.9753
7.1333 | -86.9757
7.1336 | -86.9765
7.134 | -86.9768
7.1344 | -86.9777
7.1347 | -86.9786
7.1351 | -86.9797
7.1354 | -86.9805
7.1358 | -86.9808
7.1362 | -86.9809
7.1365 | -86.9814
7.1369 | -86.9821
7.1372 | -86.9828
7.1376 | -86.9837
7.138 | -86.9848
7.1383 | -86.9862
7.1387 | -86.988
7.1391 | -86.9894
7.1394 | -86.9901
7.1398 | -86.9914
7.1401 | -86.9925
7.1405 | -86.9935
7.1409 | -86.9943
7.1412 | -86.9951
7.1416 | -86.9962
7.142 | -86.9975
7.1423 | -86.9993
7.1427 | -86.9998
7.143 | -86.9998
7.1434 | -87.0004
7.1438 | -87.0015
7.1441 | -87.0021
7.1445 | -87.0041
7.1448 | -87.0057
7.1452 | -87.0063
7.1456 | -87.0073
7.1459 | -87.0079
7.1463 | -87.008
7.1467 | -87.0092
7.147 | -87.011
7.1474 | -87.0118
7.1477 | -87.0118
7.1481 | -87.0123
7.1485 | -87.0141
7.1488 | -87.0147
7.1492 | -87.0158
7.1495 | -87.0169
7.1499 | -87.0186
7.1503 | -87.0196
7.1506 | -87.0211
7.151 | -87.0226
7.1514 | -87.0236
7.1517 | -87.0243
7.1521 | -87.025
7.1524 | -87.0252
7.1528 | -87.0256
7.1532 | -87.0265
7.1535 | -87.0271
7.1539 | -87.0285
7.1543 | -87.0299
7.1546 | -87.0309
7.155 | -87.0319
7.1553 | -87.0328
7.1557 | -87.0339
7.1561 | -87.035
7.1564 | -87.0357
7.1568 | -87.0366
7.1571 | -87.0375
7.1575 | -87.0391
7.1579 | -87.0402
7.1582 | -87.0407
7.1586 | -87.0412
7.159 | -87.0416
7.1593 | -87.0417
7.1597 | -87.0428
7.16 | -87.0438
7.1604 | -87.0452
7.1608 | -87.0457
7.1611 | -87.0464
7.1615 | -87.0459
7.1619 | -87.0439
7.1622 | -87.0439
7.1626 | -87.0449
7.1629 | -87.0454
7.1633 | -87.0457
7.1637 | -87.0468
7.164 | -87.0487
7.1644 | -87.05
7.1647 | -87.0504
7.1651 | -87.0509
7.1655 | -87.0517
7.1658 | -87.0528
7.1662 | -87.0538
7.1666 | -87.0549
7.1669 | -87.0555
7.1673 | -87.0565
7.1676 | -87.0573
7.168 | -87.0579
7.1684 | -87.0586
7.1687 | -87.0591
7.1691 | -87.0594
7.1695 | -87.0592
7.1698 | -87.0591
7.1702 | -87.0601
7.1705 | -87.0612
7.1709 | -87.0619
7.1713 | -87.0633
7.1716 | -87.0637
7.172 | -87.0642
7.1723 | -87.0648
7.1727 | -87.0656
7.1731 | -87.0665
7.1734 | -87.0661
7.1738 | -87.0661
7.1742 | -87.0688
7.1745 | -87.071
7.1749 | -87.0724
7.1752 | -87.0729
7.1756 | -87.0744
7.176 | -87.0748
7.1763 | -87.075
7.1767 | -87.075
7.177 | -87.075
7.1774 | -87.0756
7.1778 | -87.0772
7.1781 | -87.0804
7.1785 | -87.0833
7.1789 | -87.0847
7.1792 | -87.0852
7.1796 | -87.0855
7.1799 | -87.0858
7.1803 | -87.0858
7.1807 | -87.0857
7.181 | -87.0864
7.1814 | -87.0868
7.1818 | -87.088
7.1821 | -87.089
7.1825 | -87.0894
7.1828 | -87.0907
7.1832 | -87.0915
7.1836 | -87.092
7.1839 | -87.0929
7.1843 | -87.0935
7.1846 | -87.0941
7.185 | -87.0948
7.1854 | -87.0956
7.1857 | -87.0968
7.1861 | -87.098
7.1865 | -87.0982
7.1868 | -87.0982
7.1872 | -87.0981
7.1875 | -87.0982
7.1879 | -87.0995
7.1883 | -87.1007
7.1886 | -87.1012
7.189 | -87.1021
7.1893 | -87.1038
7.1897 | -87.1049
7.1901 | -87.1057
7.1904 | -87.1062
7.1908 | -87.1068
7.1912 | -87.1076
7.1915 | -87.1081
7.1919 | -87.1083
7.1922 | -87.1094
7.1926 | -87.1104
7.193 | -87.1105
7.1933 | -87.1104
7.1937 | -87.1113
7.1941 | -87.1123
7.1944 | -87.1118
7.1948 | -87.1112
7.1951 | -87.1112
7.1955 | -87.112
7.1959 | -87.1126
7.1962 | -87.1138
7.1966 | -87.1148
7.1969 | -87.1156
7.1973 | -87.1167
7.1977 | -87.1177
7.198 | -87.1184
7.1984 | -87.1184
7.1988 | -87.1184
7.1991 | -87.1193
7.1995 | -87.122
7.1998 | -87.1238
7.2002 | -87.1239
7.2006 | -87.1234
7.2009 | -87.1228
7.2013 | -87.1216
7.2016 | -87.121
7.202 | -87.1214
7.2024 | -87.1227
7.2027 | -87.1235
7.2031 | -87.124
7.2035 | -87.1247
7.2038 | -87.1254
7.2042 | -87.126
7.2045 | -87.1267
7.2049 | -87.1272
7.2053 | -87.1276
7.2056 | -87.1283
7.206 | -87.1286
7.2064 | -87.1291
7.2067 | -87.1298
7.2071 | -87.1305
7.2074 | -87.1308
7.2078 | -87.1312
7.2082 | -87.1319
7.2085 | -87.1322
7.2089 | -87.132
7.2092 | -87.1329
7.2096 | -87.1342
7.21 | -87.1356
7.2103 | -87.1361
7.2107 | -87.1371
7.2111 | -87.1378
7.2114 | -87.1384
7.2118 | -87.1389
7.2121 | -87.1414
7.2125 | -87.1428
7.2129 | -87.1433
7.2132 | -87.1437
7.2136 | -87.1445
7.214 | -87.1442
7.2143 | -87.143
7.2147 | -87.1432
7.215 | -87.1437
7.2154 | -87.1445
7.2158 | -87.1452
7.2161 | -87.1459
7.2165 | -87.1464
7.2168 | -87.1469
7.2172 | -87.1472
7.2176 | -87.1477
7.2179 | -87.1482
7.2183 | -87.1487
7.2187 | -87.1491
7.219 | -87.1491
7.2194 | -87.1495
7.2197 | -87.1504
7.2201 | -87.1504
7.2205 | -87.1501
7.2208 | -87.1504
7.2212 | -87.1507
7.2215 | -87.1511
7.2219 | -87.1517
7.2223 | -87.1521
7.2226 | -87.1523
7.223 | -87.1529
7.2234 | -87.1538
7.2237 | -87.1546
7.2241 | -87.1548
7.2244 | -87.1556
7.2248 | -87.1568
7.2252 | -87.1566
7.2255 | -87.156
7.2259 | -87.1561
7.2263 | -87.1568
7.2266 | -87.1571
7.227 | -87.1574
7.2273 | -87.158
7.2277 | -87.1582
7.2281 | -87.1583
7.2284 | -87.1585
7.2288 | -87.1592
7.2291 | -87.16
7.2295 | -87.1602
7.2299 | -87.1607
7.2302 | -87.1607
7.2306 | -87.1608
7.231 | -87.1609
7.2313 | -87.1606
7.2317 | -87.1615
7.232 | -87.1624
7.2324 | -87.1628
7.2328 | -87.1638
7.2331 | -87.1655
7.2335 | -87.1665
7.2339 | -87.1672
7.2342 | -87.1678
7.2346 | -87.1684
7.2349 | -87.1694
7.2353 | -87.17
7.2357 | -87.1707
7.236 | -87.1715
7.2364 | -87.1715
7.2367 | -87.1717
7.2371 | -87.1719
7.2375 | -87.1719
7.2378 | -87.1727
7.2382 | -87.1736
7.2386 | -87.1737
7.2389 | -87.1738
7.2393 | -87.174
7.2396 | -87.1751
7.24 | -87.176
7.2404 | -87.1766
7.2407 | -87.1778
7.2411 | -87.1785
7.2414 | -87.1789
7.2418 | -87.1796
7.2422 | -87.1796
7.2425 | -87.1804
7.2429 | -87.1806
7.2433 | -87.1811
7.2436 | -87.1815
7.244 | -87.1814
7.2443 | -87.1818
7.2447 | -87.1823
7.2451 | -87.1831
7.2454 | -87.1836
7.2458 | -87.1837
7.2462 | -87.184
7.2465 | -87.1847
7.2469 | -87.1852
7.2472 | -87.1854
7.2476 | -87.1857
7.248 | -87.1863
7.2483 | -87.1873
7.2487 | -87.1878
7.249 | -87.1884
7.2494 | -87.1891
7.2498 | -87.1894
7.2501 | -87.1897
7.2505 | -87.1907
7.2509 | -87.1919
7.2512 | -87.1931
7.2516 | -87.1936
7.2519 | -87.1939
7.2523 | -87.1951
7.2527 | -87.1959
7.253 | -87.1961
7.2534 | -87.1963
7.2537 | -87.1972
7.2541 | -87.1981
7.2545 | -87.1978
7.2548 | -87.1972
7.2552 | -87.1975
7.2556 | -87.1975
7.2559 | -87.1971
7.2563 | -87.1974
7.2566 | -87.1978
7.257 | -87.1984
7.2574 | -87.1985
7.2577 | -87.198
7.2581 | -87.1979
7.2585 | -87.1976
7.2588 | -87.1979
7.2592 | -87.1983
7.2595 | -87.199
7.2599 | -87.1997
7.2603 | -87.2006
7.2606 | -87.2009
7.261 | -87.2011
7.2613 | -87.2015
7.2617 | -87.2015
7.2621 | -87.2014
7.2624 | -87.2017
7.2628 | -87.2018
7.2632 | -87.2018
7.2635 | -87.2024
7.2639 | -87.2031
7.2642 | -87.2033
7.2646 | -87.2032
7.265 | -87.2032
7.2653 | -87.2037
7.2657 | -87.204
7.266 | -87.2035
7.2664 | -87.203
7.2668 | -87.2034
7.2671 | -87.2031
7.2675 | -87.2026
7.2679 | -87.2027
7.2682 | -87.2029
7.2686 | -87.2042
7.2689 | -87.2047
7.2693 | -87.2053
7.2697 | -87.206
7.27 | -87.2062
7.2704 | -87.2072
7.2708 | -87.2081
7.2711 | -87.2083
7.2715 | -87.2087
7.2718 | -87.2097
7.2722 | -87.2098
7.2726 | -87.2087
7.2729 | -87.2087
7.2733 | -87.2088
7.2736 | -87.2093
7.274 | -87.2105
7.2744 | -87.2113
7.2747 | -87.2109
7.2751 | -87.211
7.2755 | -87.2118
7.2758 | -87.2113
7.2762 | -87.2118
7.2765 | -87.2122
7.2769 | -87.2127
7.2773 | -87.2129
7.2776 | -87.2129
7.278 | -87.213
7.2784 | -87.2133
7.2787 | -87.2134
7.2791 | -87.2135
7.2794 | -87.2137
7.2798 | -87.2139
7.2802 | -87.2138
7.2805 | -87.2136
7.2809 | -87.214
7.2812 | -87.2145
7.2816 | -87.2143
7.282 | -87.2138
7.2823 | -87.2134
7.2827 | -87.2137
7.2831 | -87.2151
7.2834 | -87.2162
7.2838 | -87.2161
7.2841 | -87.2163
7.2845 | -87.2161
7.2849 | -87.2158
7.2852 | -87.216
7.2856 | -87.2158
7.2859 | -87.2156
7.2863 | -87.2121
7.2867 | -87.2116
7.287 | -87.2119
7.2874 | -87.2117
7.2878 | -87.2118
7.2881 | -87.2122
7.2885 | -87.2121
7.2888 | -87.212
7.2892 | -87.213
7.2896 | -87.2141
7.2899 | -87.2136
7.2903 | -87.2133
7.2907 | -87.2134
7.291 | -87.2125
7.2914 | -87.2126
7.2917 | -87.2129
7.2921 | -87.2129
7.2925 | -87.2129
7.2928 | -87.2129
7.2932 | -87.2129
7.2935 | -87.2128
7.2939 | -87.2133
7.2943 | -87.2135
7.2946 | -87.2132
7.295 | -87.2136
7.2954 | -87.2137
7.2957 | -87.2139
7.2961 | -87.2147
7.2964 | -87.2152
7.2968 | -87.2157
7.2972 | -87.2153
7.2975 | -87.2149
7.2979 | -87.2149
7.2983 | -87.2158
7.2986 | -87.2169
7.299 | -87.217
7.2993 | -87.2169
7.2997 | -87.2171
7.3001 | -87.2168
7.3004 | -87.2168
7.3008 | -87.217
7.3011 | -87.2165
7.3015 | -87.2169
7.3019 | -87.2172
7.3022 | -87.2176
7.3026 | -87.2179
7.303 | -87.2168
7.3033 | -87.2154
7.3037 | -87.2152
7.304 | -87.2156
7.3044 | -87.2155
7.3048 | -87.2148
7.3051 | -87.2145
7.3055 | -87.2142
7.3058 | -87.215
7.3062 | -87.2142
7.3066 | -87.2139
7.3069 | -87.2141
7.3073 | -87.2148
7.3077 | -87.2142
7.308 | -87.2139
7.3084 | -87.2141
7.3087 | -87.2141
7.3091 | -87.2141
7.3095 | -87.2137
7.3098 | -87.2139
7.3102 | -87.2142
7.3106 | -87.2146
7.3109 | -87.2147
7.3113 | -87.2144
7.3116 | -87.2143
7.312 | -87.2144
7.3124 | -87.2136
7.3127 | -87.2128
7.3131 | -87.2123
7.3134 | -87.212
7.3138 | -87.2119
7.3142 | -87.2121
7.3145 | -87.2112
7.3149 | -87.21
7.3153 | -87.2095
7.3156 | -87.2093
7.316 | -87.2092
7.3163 | -87.2095
7.3167 | -87.2103
7.3171 | -87.2104
7.3174 | -87.2104
7.3178 | -87.2101
7.3182 | -87.21
7.3185 | -87.2097
7.3189 | -87.2094
7.3192 | -87.2093
7.3196 | -87.2092
7.32 | -87.2096
7.3203 | -87.2094
7.3207 | -87.2092
7.321 | -87.2101
7.3214 | -87.2105
7.3218 | -87.2106
7.3221 | -87.2106
7.3225 | -87.2106
7.3229 | -87.2108
7.3232 | -87.2111
7.3236 | -87.2113
7.3239 | -87.2113
7.3243 | -87.2112
7.3247 | -87.2118
7.325 | -87.2118
7.3254 | -87.2117
7.3257 | -87.2115
7.3261 | -87.2111
7.3265 | -87.2122
7.3268 | -87.2129
7.3272 | -87.213
7.3276 | -87.2138
7.3279 | -87.2139
7.3283 | -87.2144
7.3286 | -87.2139
7.329 | -87.2126
7.3294 | -87.2117
7.3297 | -87.2123
7.3301 | -87.2123
7.3305 | -87.2121
7.3308 | -87.2118
7.3312 | -87.2114
7.3315 | -87.2106
7.3319 | -87.2099
7.3323 | -87.2099
7.3326 | -87.2095
7.333 | -87.2088
7.3333 | -87.2084
7.3337 | -87.2081
7.3341 | -87.208
7.3344 | -87.2077
7.3348 | -87.2074
7.3352 | -87.2068
7.3355 | -87.2063
7.3359 | -87.2071
7.3362 | -87.2083
7.3366 | -87.2085
7.337 | -87.2078
7.3373 | -87.2081
7.3377 | -87.208
7.338 | -87.2079
7.3384 | -87.2076
7.3388 | -87.2083
7.3391 | -87.2086
7.3395 | -87.2079
7.3399 | -87.2073
7.3402 | -87.2056
7.3406 | -87.2048
7.3409 | -87.2046
7.3413 | -87.2034
7.3417 | -87.2023
7.342 | -87.202
7.3424 | -87.2022
7.3428 | -87.2023
7.3431 | -87.2014
7.3435 | -87.2007
7.3438 | -87.2004
7.3442 | -87.2001
7.3446 | -87.1997
7.3449 | -87.1992
7.3453 | -87.1986
7.3456 | -87.1983
7.346 | -87.1983
7.3464 | -87.1983
7.3467 | -87.1981
7.3471 | -87.198
7.3475 | -87.1973
7.3478 | -87.1963
7.3482 | -87.1952
7.3485 | -87.1949
7.3489 | -87.1936
7.3493 | -87.1935
7.3496 | -87.1937
7.35 | -87.1938
7.3503 | -87.1933
7.3507 | -87.1929
7.3511 | -87.1928
7.3514 | -87.193
7.3518 | -87.1943
7.3522 | -87.1943
7.3525 | -87.1936
7.3529 | -87.1935
7.3532 | -87.194
7.3536 | -87.1935
7.354 | -87.1922
7.3543 | -87.1906
7.3547 | -87.1898
7.3551 | -87.1895
7.3554 | -87.1893
7.3558 | -87.1893
7.3561 | -87.189
7.3565 | -87.188
7.3569 | -87.1872
7.3572 | -87.187
7.3576 | -87.1866
7.3579 | -87.1858
7.3583 | -87.1849
7.3587 | -87.1844
7.359 | -87.1833
7.3594 | -87.1831
7.3598 | -87.1829
7.3601 | -87.1828
7.3605 | -87.1822
7.3608 | -87.1812
7.3612 | -87.1809
7.3616 | -87.1802
7.3619 | -87.179
7.3623 | -87.1789
7.3627 | -87.1781
7.363 | -87.177
7.3634 | -87.176
7.3637 | -87.1749
7.3641 | -87.1742
7.3645 | -87.1737
7.3648 | -87.1734
7.3652 | -87.1728
7.3655 | -87.1716
7.3659 | -87.171
7.3663 | -87.1706
7.3666 | -87.17
7.367 | -87.1693
7.3674 | -87.1691
7.3677 | -87.1685
7.3681 | -87.1691
7.3684 | -87.1692
7.3688 | -87.1694
7.3692 | -87.1698
7.3695 | -87.1688
7.3699 | -87.1676
7.3702 | -87.1666
7.3706 | -87.1656
7.371 | -87.1656
7.3713 | -87.1652
7.3717 | -87.1647
7.3721 | -87.1643
7.3724 | -87.1633
7.3728 | -87.1625
7.3731 | -87.162
7.3735 | -87.1616
7.3739 | -87.1606
7.3742 | -87.1596
7.3746 | -87.1591
7.375 | -87.1584
7.3753 | -87.1579
7.3757 | -87.1571
7.376 | -87.1568
7.3764 | -87.1569
7.3768 | -87.156
7.3771 | -87.1553
7.3775 | -87.1551
7.3778 | -87.1542
7.3782 | -87.1537
7.3786 | -87.1532
7.3789 | -87.1526
7.3793 | -87.1522
7.3797 | -87.1522
7.38 | -87.1519
7.3804 | -87.1512
7.3807 | -87.1507
7.3811 | -87.1504
7.3815 | -87.1504
7.3818 | -87.1497
7.3822 | -87.1492
7.3826 | -87.1484
7.3829 | -87.1476
7.3833 | -87.1469
7.3836 | -87.1483
7.384 | -87.1484
7.3844 | -87.1489
7.3847 | -87.149
7.3851 | -87.149
7.3854 | -87.1485
7.3858 | -87.1479
7.3862 | -87.1473
7.3865 | -87.1459
7.3869 | -87.1453
7.3873 | -87.1458
7.3876 | -87.1451
7.388 | -87.1447
7.3883 | -87.144
7.3887 | -87.1433
7.3891 | -87.1428
7.3894 | -87.142
7.3898 | -87.1411
7.3901 | -87.1405
7.3905 | -87.1403
7.3909 | -87.1399
7.3912 | -87.1393
7.3916 | -87.1383
7.392 | -87.1372
7.3923 | -87.1362
7.3927 | -87.1351
7.393 | -87.134
7.3934 | -87.1335
7.3938 | -87.1328
7.3941 | -87.1323
7.3945 | -87.1315
7.3949 | -87.1306
7.3952 | -87.1306
7.3956 | -87.131
7.3959 | -87.1308
7.3963 | -87.1303
7.3967 | -87.1295
7.397 | -87.1286
7.3974 | -87.1279
7.3977 | -87.1275
7.3981 | -87.1269
7.3985 | -87.1264
7.3988 | -87.126
7.3992 | -87.1255
7.3996 | -87.125
7.3999 | -87.1253
7.4003 | -87.1255
7.4006 | -87.125
7.401 | -87.1237
7.4014 | -87.1227
7.4017 | -87.1222
7.4021 | -87.1216
7.4024 | -87.1202
7.4028 | -87.1191
7.4032 | -87.1186
7.4035 | -87.1178
7.4039 | -87.117
7.4043 | -87.1164
7.4046 | -87.1171
7.405 | -87.1177
7.4053 | -87.1179
7.4057 | -87.1184
7.4061 | -87.1186
7.4064 | -87.1183
7.4068 | -87.1178
7.4072 | -87.117
7.4075 | -87.1167
7.4079 | -87.1172
7.4082 | -87.1176
7.4086 | -87.1174
7.409 | -87.1175
7.4093 | -87.1183
7.4097 | -87.1188
7.41 | -87.1187
7.4104 | -87.1183
7.4108 | -87.1178
7.4111 | -87.1172
7.4115 | -87.1172
7.4119 | -87.1164
7.4122 | -87.1155
7.4126 | -87.115
7.4129 | -87.1139
7.4133 | -87.113
7.4137 | -87.1123
7.414 | -87.1118
7.4144 | -87.1112
7.4148 | -87.1106
7.4151 | -87.1101
7.4155 | -87.1094
7.4158 | -87.1087
7.4162 | -87.108
7.4166 | -87.1069
7.4169 | -87.106
7.4173 | -87.1054
7.4176 | -87.1046
7.418 | -87.1034
7.4184 | -87.1027
7.4187 | -87.1022
7.4191 | -87.1011
7.4195 | -87.1005
7.4198 | -87.0999
7.4202 | -87.0992
7.4205 | -87.0983
7.4209 | -87.0976
7.4213 | -87.0968
7.4216 | -87.0962
7.422 | -87.0962
7.4223 | -87.0956
7.4227 | -87.0951
7.4231 | -87.0947
7.4234 | -87.0942
7.4238 | -87.0944
7.4242 | -87.0949
7.4245 | -87.0947
7.4249 | -87.0941
7.4252 | -87.0931
7.4256 | -87.0924
7.426 | -87.0921
7.4263 | -87.0915
7.4267 | -87.0915
7.4271 | -87.0918
7.4274 | -87.0923
7.4278 | -87.092
7.4281 | -87.0918
7.4285 | -87.0917
7.4289 | -87.092
7.4292 | -87.0919
7.4296 | -87.0908
7.4299 | -87.0901
7.4303 | -87.0901
7.4307 | -87.0906
7.431 | -87.0905
7.4314 | -87.0901
7.4318 | -87.0898
7.4321 | -87.0894
7.4325 | -87.0879
7.4328 | -87.0864
7.4332 | -87.0859
7.4336 | -87.0856
7.4339 | -87.0849
7.4343 | -87.0838
7.4347 | -87.0831
7.435 | -87.0827
7.4354 | -87.0818
7.4357 | -87.0806
7.4361 | -87.0805
7.4365 | -87.0806
7.4368 | -87.0807
7.4372 | -87.0809
7.4375 | -87.0807
7.4379 | -87.0807
7.4383 | -87.0816
7.4386 | -87.0825
7.439 | -87.0827
7.4394 | -87.0834
7.4397 | -87.0853
7.4401 | -87.0858
7.4404 | -87.0853
7.4408 | -87.0854
7.4412 | -87.0856
7.4415 | -87.0855
7.4419 | -87.0849
7.4422 | -87.084
7.4426 | -87.0838
7.443 | -87.0838
7.4433 | -87.0834
7.4437 | -87.0829
7.4441 | -87.0823
7.4444 | -87.0819
7.4448 | -87.0816
7.4451 | -87.0813
7.4455 | -87.0814
7.4459 | -87.081
7.4462 | -87.081
7.4466 | -87.081
7.447 | -87.0812
7.4473 | -87.083
7.4477 | -87.0843
7.448 | -87.0845
7.4484 | -87.0844
7.4488 | -87.0842
7.4491 | -87.0833
7.4495 | -87.0824
7.4498 | -87.0821
7.4502 | -87.0821
7.4506 | -87.0825
7.4509 | -87.0842
7.4513 | -87.0847
7.4517 | -87.084
7.452 | -87.0845
7.4524 | -87.0851
7.4527 | -87.0853
7.4531 | -87.0849
7.4535 | -87.0843
7.4538 | -87.0831
7.4542 | -87.0816
7.4545 | -87.081
7.4549 | -87.0808
7.4553 | -87.0807
7.4556 | -87.0807
7.456 | -87.0806
7.4564 | -87.0804
7.4567 | -87.08
7.4571 | -87.0797
7.4574 | -87.079
7.4578 | -87.0787
7.4582 | -87.0784
7.4585 | -87.0781
7.4589 | -87.0779
7.4593 | -87.0774
7.4596 | -87.0768
7.46 | -87.0761
7.4603 | -87.0757
7.4607 | -87.0753
7.4611 | -87.0751
7.4614 | -87.0747
7.4618 | -87.0744
7.4621 | -87.0743
7.4625 | -87.0743
7.4629 | -87.0742
7.4632 | -87.0747
7.4636 | -87.0748
7.464 | -87.0748
7.4643 | -87.0743
7.4647 | -87.074
7.465 | -87.0731
7.4654 | -87.0718
7.4658 | -87.0708
7.4661 | -87.07
7.4665 | -87.0695
7.4668 | -87.0698
7.4672 | -87.0703
7.4676 | -87.0714
7.4679 | -87.072
7.4683 | -87.0724
7.4687 | -87.0727
7.469 | -87.0729
7.4694 | -87.0727
7.4697 | -87.0726
7.4701 | -87.0723
7.4705 | -87.0723
7.4708 | -87.0721
7.4712 | -87.0718
7.4716 | -87.0716
7.4719 | -87.0716
7.4723 | -87.0715
7.4726 | -87.0718
7.473 | -87.072
7.4734 | -87.0718
7.4737 | -87.0717
7.4741 | -87.0715
7.4744 | -87.0709
7.4748 | -87.0703
7.4752 | -87.0694
7.4755 | -87.069
7.4759 | -87.0691
7.4763 | -87.0689
7.4766 | -87.0688
7.477 | -87.0687
7.4773 | -87.0682
7.4777 | -87.0678
7.4781 | -87.0677
7.4784 | -87.0676
7.4788 | -87.0679
7.4792 | -87.0679
7.4795 | -87.0672
7.4799 | -87.0672
7.4802 | -87.0672
7.4806 | -87.0672
7.481 | -87.0664
7.4813 | -87.0654
7.4817 | -87.0647
7.482 | -87.0641
7.4824 | -87.0642
7.4828 | -87.064
7.4831 | -87.0639
7.4835 | -87.0648
7.4839 | -87.0652
7.4842 | -87.0656
7.4846 | -87.066
7.4849 | -87.0663
7.4853 | -87.0663
7.4857 | -87.0669
7.486 | -87.0685
7.4864 | -87.069
7.4867 | -87.0691
7.4871 | -87.0689
7.4875 | -87.0681
7.4878 | -87.068
7.4882 | -87.0683
7.4886 | -87.0686
7.4889 | -87.0689
7.4893 | -87.0685
7.4896 | -87.0684
7.49 | -87.0687
7.4904 | -87.0683
7.4907 | -87.0685
7.4911 | -87.0688
7.4915 | -87.0691
7.4918 | -87.0692
7.4922 | -87.0692
7.4925 | -87.0688
7.4929 | -87.0686
7.4933 | -87.0694
7.4936 | -87.07
7.494 | -87.0701
7.4943 | -87.0696
7.4947 | -87.069
7.4951 | -87.0699
7.4954 | -87.0706
7.4958 | -87.0716
7.4962 | -87.0723
7.4965 | -87.0732
7.4969 | -87.073
7.4972 | -87.0723
7.4976 | -87.0722
7.498 | -87.072
7.4983 | -87.0721
7.4987 | -87.0724
7.4991 | -87.0726
7.4994 | -87.0724
7.4998 | -87.073
7.5001 | -87.0737
7.5005 | -87.0738
7.5009 | -87.0741
7.5012 | -87.0743
7.5016 | -87.0736
7.5019 | -87.0735
7.5023 | -87.0745
7.5027 | -87.0751
7.503 | -87.0753
7.5034 | -87.0755
7.5038 | -87.076
7.5041 | -87.0763
7.5045 | -87.0758
7.5048 | -87.0758
7.5052 | -87.0761
7.5056 | -87.0756
7.5059 | -87.0755
7.5063 | -87.0757
7.5066 | -87.0758
7.507 | -87.0765
7.5074 | -87.0773
7.5077 | -87.0775
7.5081 | -87.0778
7.5085 | -87.0779
7.5088 | -87.0783
7.5092 | -87.0785
7.5095 | -87.0788
7.5099 | -87.0794
7.5103 | -87.0803
7.5106 | -87.0809
7.511 | -87.0812
7.5114 | -87.0805
7.5117 | -87.0798
7.5121 | -87.0797
7.5124 | -87.0803
7.5128 | -87.081
7.5132 | -87.0815
7.5135 | -87.0819
7.5139 | -87.082
7.5142 | -87.0824
7.5146 | -87.0829
7.515 | -87.0823
7.5153 | -87.0808
7.5157 | -87.0819
7.5161 | -87.0822
7.5164 | -87.0823
7.5168 | -87.0824
7.5171 | -87.0827
7.5175 | -87.0834
7.5179 | -87.0847
7.5182 | -87.0857
7.5186 | -87.0862
7.5189 | -87.0862
7.5193 | -87.087
7.5197 | -87.088
7.52 | -87.0882
7.5204 | -87.0889
7.5208 | -87.0894
7.5211 | -87.0901
7.5215 | -87.0904
7.5218 | -87.0909
7.5222 | -87.0923
7.5226 | -87.0932
7.5229 | -87.0937
7.5233 | -87.0945
7.5237 | -87.095
7.524 | -87.0939
7.5244 | -87.0938
7.5247 | -87.0945
7.5251 | -87.0947
7.5255 | -87.0951
7.5258 | -87.0957
7.5262 | -87.0967
7.5265 | -87.0977
7.5269 | -87.0981
7.5273 | -87.0986
7.5276 | -87.0991
7.528 | -87.0988
7.5284 | -87.0986
7.5287 | -87.0997
7.5291 | -87.1014
7.5294 | -87.1027
7.5298 | -87.1035
7.5302 | -87.1039
7.5305 | -87.1044
7.5309 | -87.1048
7.5313 | -87.106
7.5316 | -87.1065
7.532 | -87.1075
7.5323 | -87.1079
7.5327 | -87.1083
7.5331 | -87.1088
7.5334 | -87.1092
7.5338 | -87.1096
7.5341 | -87.1101
7.5345 | -87.1108
7.5349 | -87.1113
7.5352 | -87.1115
7.5356 | -87.1118
7.536 | -87.1124
7.5363 | -87.1127
7.5367 | -87.1139
7.537 | -87.1148
7.5374 | -87.1157
7.5378 | -87.1166
7.5381 | -87.1175
7.5385 | -87.1182
7.5388 | -87.119
7.5392 | -87.1199
7.5396 | -87.1207
7.5399 | -87.1215
7.5403 | -87.1221
7.5407 | -87.1226
7.541 | -87.1232
7.5414 | -87.1237
7.5417 | -87.1242
7.5421 | -87.1249
7.5425 | -87.1256
7.5428 | -87.1259
7.5432 | -87.1263
7.5436 | -87.1266
7.5439 | -87.1266
7.5443 | -87.1275
7.5446 | -87.128
7.545 | -87.1284
7.5454 | -87.1293
7.5457 | -87.1307
7.5461 | -87.1313
7.5464 | -87.1319
7.5468 | -87.1323
7.5472 | -87.134
7.5475 | -87.1342
7.5479 | -87.1346
7.5483 | -87.135
7.5486 | -87.1357
7.549 | -87.1371
7.5493 | -87.1384
7.5497 | -87.139
7.5501 | -87.1402
7.5504 | -87.1414
7.5508 | -87.1421
7.5511 | -87.1426
7.5515 | -87.1432
7.5519 | -87.1439
7.5522 | -87.1446
7.5526 | -87.1453
7.553 | -87.146
7.5533 | -87.1464
7.5537 | -87.1463
7.554 | -87.1464
7.5544 | -87.147
7.5548 | -87.1478
7.5551 | -87.1485
7.5555 | -87.149
7.5559 | -87.1495
7.5562 | -87.1503
7.5566 | -87.1511
7.5569 | -87.1525
7.5573 | -87.154
7.5577 | -87.1554
7.558 | -87.1564
7.5584 | -87.1568
7.5587 | -87.1574
7.5591 | -87.1586
7.5595 | -87.1594
7.5598 | -87.1604
7.5602 | -87.1614
7.5606 | -87.1622
7.5609 | -87.1632
7.5613 | -87.1639
7.5616 | -87.1646
7.562 | -87.1656
7.5624 | -87.1663
7.5627 | -87.1684
7.5631 | -87.1706
7.5635 | -87.1714
7.5638 | -87.1723
7.5642 | -87.1734
7.5645 | -87.174
7.5649 | -87.1749
7.5653 | -87.1762
7.5656 | -87.1773
7.566 | -87.1781
7.5663 | -87.1788
7.5667 | -87.1797
7.5671 | -87.181
7.5674 | -87.1823
7.5678 | -87.1834
7.5682 | -87.185
7.5685 | -87.1856
7.5689 | -87.1863
7.5692 | -87.1872
7.5696 | -87.1885
7.57 | -87.19
7.5703 | -87.1913
7.5707 | -87.1925
7.571 | -87.1939
7.5714 | -87.1954
7.5718 | -87.1971
7.5721 | -87.1976
7.5725 | -87.1988
7.5729 | -87.1995
7.5732 | -87.2004
7.5736 | -87.2014
7.5739 | -87.2021
7.5743 | -87.203
7.5747 | -87.2039
7.575 | -87.205
7.5754 | -87.2053
7.5758 | -87.2051
7.5761 | -87.2057
7.5765 | -87.2063
7.5768 | -87.2073
7.5772 | -87.2081
7.5776 | -87.2091
7.5779 | -87.2101
7.5783 | -87.2111
7.5786 | -87.2122
7.579 | -87.2131
7.5794 | -87.214
7.5797 | -87.2151
7.5801 | -87.2169
7.5805 | -87.2186
7.5808 | -87.2198
7.5812 | -87.2211
7.5815 | -87.2228
7.5819 | -87.2245
7.5823 | -87.2256
7.5826 | -87.2268
7.583 | -87.2281
7.5834 | -87.2291
7.5837 | -87.2299
7.5841 | -87.2308
7.5844 | -87.2315
7.5848 | -87.2319
7.5852 | -87.2323
7.5855 | -87.2318
7.5859 | -87.2311
7.5862 | -87.2309
7.5866 | -87.2306
7.587 | -87.2313
7.5873 | -87.2326
7.5877 | -87.2336
7.5881 | -87.2347
7.5884 | -87.2358
7.5888 | -87.2376
7.5891 | -87.2387
7.5895 | -87.2397
7.5899 | -87.2413
7.5902 | -87.2427
7.5906 | -87.2448
7.5909 | -87.2465
7.5913 | -87.248
7.5917 | -87.2496
7.592 | -87.2509
7.5924 | -87.2519
7.5928 | -87.2526
7.5931 | -87.2529
7.5935 | -87.2534
7.5938 | -87.2543
7.5942 | -87.255
7.5946 | -87.256
7.5949 | -87.2572
7.5953 | -87.2584
7.5957 | -87.2597
7.596 | -87.2612
7.5964 | -87.2621
7.5967 | -87.2633
7.5971 | -87.2643
7.5975 | -87.2655
7.5978 | -87.2668
7.5982 | -87.2681
7.5985 | -87.269
7.5989 | -87.2698
7.5993 | -87.2711
7.5996 | -87.2721
7.6 | -87.2734
7.6004 | -87.2746
7.6007 | -87.2758
7.6011 | -87.2773
7.6014 | -87.2786
7.6018 | -87.2797
7.6022 | -87.2809
7.6025 | -87.2825
7.6029 | -87.2841
7.6032 | -87.2855
7.6036 | -87.2864
7.604 | -87.2875
7.6043 | -87.2889
7.6047 | -87.2899
7.6051 | -87.2909
7.6054 | -87.2921
7.6058 | -87.2935
7.6061 | -87.2956
7.6065 | -87.297
7.6069 | -87.2982
7.6072 | -87.2994
7.6076 | -87.3006
7.608 | -87.3024
7.6083 | -87.3041
7.6087 | -87.3057
7.609 | -87.3074
7.6094 | -87.3087
7.6098 | -87.3102
7.6101 | -87.3115
7.6105 | -87.3127
7.6108 | -87.3136
7.6112 | -87.3147
7.6116 | -87.3171
7.6119 | -87.3199
7.6123 | -87.3221
7.6127 | -87.3238
7.613 | -87.325
7.6134 | -87.3263
7.6137 | -87.3272
7.6141 | -87.3277
7.6145 | -87.3286
7.6148 | -87.3302
7.6152 | -87.3322
7.6155 | -87.3343
7.6159 | -87.3355
7.6163 | -87.3365
7.6166 | -87.3377
7.617 | -87.3392
7.6174 | -87.3406
7.6177 | -87.342
7.6181 | -87.3435
7.6184 | -87.3447
7.6188 | -87.3461
7.6192 | -87.347
7.6195 | -87.3484
7.6199 | -87.3499
7.6203 | -87.3516
7.6206 | -87.3529
7.621 | -87.3542
7.6213 | -87.3554
7.6217 | -87.3563
7.6221 | -87.3574
7.6224 | -87.3587
7.6228 | -87.3603
7.6231 | -87.3614
7.6235 | -87.3631
7.6239 | -87.3641
7.6242 | -87.3654
7.6246 | -87.3665
7.625 | -87.3682
7.6253 | -87.3697
7.6257 | -87.3709
7.626 | -87.372
7.6264 | -87.3731
7.6268 | -87.3746
7.6271 | -87.3763
7.6275 | -87.3774
7.6279 | -87.3788
7.6282 | -87.3799
7.6286 | -87.3815
7.6289 | -87.3836
7.6293 | -87.3847
7.6297 | -87.3861
7.63 | -87.3875
7.6304 | -87.389
7.6307 | -87.3906
7.6311 | -87.3926
7.6315 | -87.3944
7.6318 | -87.3957
7.6322 | -87.3972
7.6326 | -87.3986
7.6329 | -87.3997
7.6333 | -87.4009
7.6336 | -87.4024
7.634 | -87.4041
7.6344 | -87.4055
7.6347 | -87.4069
7.6351 | -87.4083
7.6354 | -87.4097
7.6358 | -87.4112
7.6362 | -87.4126
7.6365 | -87.4142
7.6369 | -87.4157
7.6373 | -87.4174
7.6376 | -87.419
7.638 | -87.4207
7.6383 | -87.4224
7.6387 | -87.4239
7.6391 | -87.426
7.6394 | -87.4277
7.6398 | -87.4292
7.6402 | -87.4309
7.6405 | -87.4326
7.6409 | -87.4343
7.6412 | -87.436
7.6416 | -87.4373
7.642 | -87.4389
7.6423 | -87.4405
7.6427 | -87.4422
7.643 | -87.4442
7.6434 | -87.4461
7.6438 | -87.4476
7.6441 | -87.4495
7.6445 | -87.4512
7.6449 | -87.4527
7.6452 | -87.4541
7.6456 | -87.4552
7.6459 | -87.4562
7.6463 | -87.4571
7.6467 | -87.4586
7.647 | -87.461
7.6474 | -87.4629
7.6478 | -87.4645
7.6481 | -87.466
7.6485 | -87.4675
7.6488 | -87.4695
7.6492 | -87.4718
7.6496 | -87.4736
7.6499 | -87.4747
7.6503 | -87.4766
7.6506 | -87.478
7.651 | -87.4793
7.6514 | -87.4811
7.6517 | -87.4825
7.6521 | -87.4832
7.6525 | -87.4842
7.6528 | -87.4862
7.6532 | -87.4881
7.6535 | -87.4892
7.6539 | -87.4904
7.6543 | -87.4916
7.6546 | -87.4927
7.655 | -87.4938
7.6553 | -87.495
7.6557 | -87.4964
7.6561 | -87.4975
7.6564 | -87.4991
7.6568 | -87.5003
7.6572 | -87.5015
7.6575 | -87.5026
7.6579 | -87.5041
7.6582 | -87.5058
7.6586 | -87.5076
7.659 | -87.5087
7.6593 | -87.5094
7.6597 | -87.5108
7.6601 | -87.5128
7.6604 | -87.5146
7.6608 | -87.5161
7.6611 | -87.518
7.6615 | -87.52
7.6619 | -87.5209
7.6622 | -87.5228
7.6626 | -87.5242
7.6629 | -87.525
7.6633 | -87.5265
7.6637 | -87.5278
7.664 | -87.5294
7.6644 | -87.531
7.6648 | -87.5326
7.6651 | -87.5337
7.6655 | -87.5351
7.6658 | -87.5365
7.6662 | -87.5371
7.6666 | -87.5384
7.6669 | -87.5404
7.6673 | -87.542
7.6677 | -87.5435
7.668 | -87.545
7.6684 | -87.5464
7.6687 | -87.548
7.6691 | -87.5491
7.6695 | -87.5507
7.6698 | -87.5525
7.6702 | -87.5539
7.6705 | -87.5554
7.6709 | -87.5569
7.6713 | -87.5581
7.6716 | -87.5593
7.672 | -87.5612
7.6724 | -87.5633
7.6727 | -87.5652
7.6731 | -87.5672
7.6734 | -87.569
7.6738 | -87.5707
7.6742 | -87.573
7.6745 | -87.5747
7.6749 | -87.5764
7.6752 | -87.578
7.6756 | -87.5793
7.676 | -87.5808
7.6763 | -87.5825
7.6767 | -87.5858
7.6771 | -87.59
7.6774 | -87.5927
7.6778 | -87.5939
7.6781 | -87.5953
7.6785 | -87.5966
7.6789 | -87.5968
7.6792 | -87.5972
7.6796 | -87.5983
7.68 | -87.5998
7.6803 | -87.6016
7.6807 | -87.6032
7.681 | -87.605
7.6814 | -87.6067
7.6818 | -87.6079
7.6821 | -87.6093
7.6825 | -87.6106
7.6828 | -87.6116
7.6832 | -87.6131
7.6836 | -87.6144
7.6839 | -87.6153
7.6843 | -87.6168
7.6847 | -87.6184
7.685 | -87.6209
7.6854 | -87.6231
7.6857 | -87.6246
7.6861 | -87.626
7.6865 | -87.6268
7.6868 | -87.628
7.6872 | -87.6293
7.6875 | -87.6308
7.6879 | -87.6329
7.6883 | -87.6346
7.6886 | -87.636
7.689 | -87.6375
7.6894 | -87.6396
7.6897 | -87.6416
7.6901 | -87.6437
7.6904 | -87.6454
7.6908 | -87.6468
7.6912 | -87.6484
7.6915 | -87.6497
7.6919 | -87.6513
7.6923 | -87.6532
7.6926 | -87.6554
7.693 | -87.6574
7.6933 | -87.6581
7.6937 | -87.6598
7.6941 | -87.6612
7.6944 | -87.6624
7.6948 | -87.6637
7.6951 | -87.6651
7.6955 | -87.6665
7.6959 | -87.6682
7.6962 | -87.6698
7.6966 | -87.6721
7.697 | -87.6743
7.6973 | -87.6756
7.6977 | -87.677
7.698 | -87.6773
7.6984 | -87.677
7.6988 | -87.6781
7.6991 | -87.679
7.6995 | -87.6801
7.6998 | -87.6816
7.7002 | -87.6831
7.7006 | -87.6848
7.7009 | -87.6865
7.7013 | -87.6881
7.7017 | -87.6896
7.702 | -87.6908
7.7024 | -87.6923
7.7027 | -87.6938
7.7031 | -87.6957
7.7035 | -87.6976
7.7038 | -87.6978
7.7042 | -87.6998
7.7046 | -87.702
7.7049 | -87.7036
7.7053 | -87.705
7.7056 | -87.7066
7.706 | -87.7081
7.7064 | -87.7096
7.7067 | -87.711
7.7071 | -87.7131
7.7074 | -87.7148
7.7078 | -87.7164
7.7082 | -87.7175
7.7085 | -87.7183
7.7089 | -87.7196
7.7093 | -87.7212
7.7096 | -87.7224
7.71 | -87.7241
7.7103 | -87.7263
7.7107 | -87.7284
7.7111 | -87.7299
7.7114 | -87.7307
7.7118 | -87.7329
7.7122 | -87.7361
7.7125 | -87.7379
7.7129 | -87.739
7.7132 | -87.7393
7.7136 | -87.7409
7.714 | -87.7426
7.7143 | -87.7443
7.7147 | -87.7457
7.715 | -87.747
7.7154 | -87.7478
7.7158 | -87.7487
7.7161 | -87.7508
7.7165 | -87.7534
7.7169 | -87.756
7.7172 | -87.7574
7.7176 | -87.7583
7.7179 | -87.7603
7.7183 | -87.7628
7.7187 | -87.7654
7.719 | -87.7687
7.7194 | -87.7713
7.7197 | -87.773
7.7201 | -87.7744
7.7205 | -87.7756
7.7208 | -87.7767
7.7212 | -87.7788
7.7216 | -87.7802
7.7219 | -87.7815
7.7223 | -87.783
7.7226 | -87.7845
7.723 | -87.7855
7.7234 | -87.7858
7.7237 | -87.7868
7.7241 | -87.7878
7.7245 | -87.7884
7.7248 | -87.7898
7.7252 | -87.7911
7.7255 | -87.7926
7.7259 | -87.7945
7.7263 | -87.7964
7.7266 | -87.7976
7.727 | -87.7996
7.7273 | -87.8013
7.7277 | -87.8022
7.7281 | -87.8035
7.7284 | -87.8052
7.7288 | -87.8073
7.7292 | -87.8084
7.7295 | -87.8098
7.7299 | -87.8109
7.7302 | -87.8124
7.7306 | -87.8141
7.731 | -87.8166
7.7313 | -87.8187
7.7317 | -87.8205
7.7321 | -87.8219
7.7324 | -87.8236
7.7328 | -87.8259
7.7331 | -87.8294
7.7335 | -87.8313
7.7339 | -87.8326
7.7342 | -87.8341
7.7346 | -87.8353
7.7349 | -87.8355
7.7353 | -87.8369
7.7357 | -87.8381
7.736 | -87.8389
7.7364 | -87.8401
7.7368 | -87.8416
7.7371 | -87.843
7.7375 | -87.8444
7.7378 | -87.8454
7.7382 | -87.8474
7.7386 | -87.8489
7.7389 | -87.8499
7.7393 | -87.8515
7.7396 | -87.8535
7.74 | -87.8551
7.7404 | -87.8566
7.7407 | -87.8585
7.7411 | -87.8595
7.7415 | -87.8604
7.7418 | -87.862
7.7422 | -87.8631
7.7425 | -87.8643
7.7429 | -87.8657
7.7433 | -87.867
7.7436 | -87.8686
7.744 | -87.8698
7.7444 | -87.8714
7.7447 | -87.874
7.7451 | -87.8765
7.7454 | -87.8773
7.7458 | -87.8789
7.7462 | -87.8802
7.7465 | -87.8815
7.7469 | -87.883
7.7472 | -87.884
7.7476 | -87.885
7.748 | -87.8862
7.7483 | -87.8877
7.7487 | -87.8902
7.7491 | -87.893
7.7494 | -87.895
7.7498 | -87.8968
7.7501 | -87.8992
7.7505 | -87.9007
7.7509 | -87.9027
7.7512 | -87.9048
7.7516 | -87.9055
7.7519 | -87.9067
7.7523 | -87.9081
7.7527 | -87.9088
7.753 | -87.9103
7.7534 | -87.9118
7.7538 | -87.9128
7.7541 | -87.9138
7.7545 | -87.9153
7.7548 | -87.9168
7.7552 | -87.9182
7.7556 | -87.9195
7.7559 | -87.9211
7.7563 | -87.9229
7.7567 | -87.9251
7.757 | -87.9269
7.7574 | -87.9282
7.7577 | -87.9296
7.7581 | -87.931
7.7585 | -87.9323
7.7588 | -87.9329
7.7592 | -87.9345
7.7595 | -87.9363
7.7599 | -87.9379
7.7603 | -87.9392
7.7606 | -87.9402
7.761 | -87.9408
7.7614 | -87.9412
7.7617 | -87.942
7.7621 | -87.943
7.7624 | -87.9441
7.7628 | -87.9442
7.7632 | -87.945
7.7635 | -87.9463
7.7639 | -87.9472
7.7642 | -87.9483
7.7646 | -87.9488
7.765 | -87.9497
7.7653 | -87.9506
7.7657 | -87.9514
7.7661 | -87.9519
7.7664 | -87.9509
7.7668 | -87.9519
7.7671 | -87.9534
7.7675 | -87.9548
7.7679 | -87.9557
7.7682 | -87.9571
7.7686 | -87.9578
7.769 | -87.9582
7.7693 | -87.9598
7.7697 | -87.9598
7.77 | -87.9593
7.7704 | -87.9601
7.7708 | -87.9604
7.7711 | -87.9607
7.7715 | -87.9613
7.7718 | -87.9623
7.7722 | -87.9641
7.7726 | -87.9654
7.7729 | -87.9668
7.7733 | -87.9687
7.7737 | -87.9703
7.774 | -87.9714
7.7744 | -87.9727
7.7747 | -87.9738
7.7751 | -87.9753
7.7755 | -87.977
7.7758 | -87.9789
7.7762 | -87.9805
7.7766 | -87.9804
7.7769 | -87.9802
7.7773 | -87.9812
7.7776 | -87.9824
7.778 | -87.9839
7.7784 | -87.9845
7.7787 | -87.9848
7.7791 | -87.9863
7.7794 | -87.9869
7.7798 | -87.9879
7.7802 | -87.9895
7.7805 | -87.9908
7.7809 | -87.9915
7.7813 | -87.9926
7.7816 | -87.9934
7.782 | -87.9944
7.7823 | -87.9944
7.7827 | -87.9951
7.7831 | -87.9962
7.7834 | -87.9973
7.7838 | -87.9989
7.7841 | -88.0006
7.7845 | -88.0016
7.7849 | -88.0029
7.7852 | -88.0043
7.7856 | -88.0044
7.786 | -88.005
7.7863 | -88.0062
7.7867 | -88.0077
7.787 | -88.0085
7.7874 | -88.0098
7.7878 | -88.0108
7.7881 | -88.0121
7.7885 | -88.0135
7.7889 | -88.0158
7.7892 | -88.0178
7.7896 | -88.019
7.7899 | -88.0209
7.7903 | -88.023
7.7907 | -88.0241
7.791 | -88.0257
7.7914 | -88.027
7.7917 | -88.0279
7.7921 | -88.0288
7.7925 | -88.0297
7.7928 | -88.0309
7.7932 | -88.0325
7.7936 | -88.0336
7.7939 | -88.0346
7.7943 | -88.0356
7.7946 | -88.0364
7.795 | -88.0381
7.7954 | -88.0405
7.7957 | -88.0422
7.7961 | -88.0433
7.7965 | -88.0441
7.7968 | -88.0444
7.7972 | -88.0452
7.7975 | -88.0454
7.7979 | -88.0464
7.7983 | -88.0466
7.7986 | -88.0473
7.799 | -88.0483
7.7993 | -88.0498
7.7997 | -88.0506
7.8001 | -88.0506
7.8004 | -88.0513
7.8008 | -88.052
7.8012 | -88.0528
7.8015 | -88.0536
7.8019 | -88.0548
7.8022 | -88.0561
7.8026 | -88.0564
7.803 | -88.056
7.8033 | -88.0567
7.8037 | -88.0576
7.804 | -88.058
7.8044 | -88.0578
7.8048 | -88.0581
7.8051 | -88.0591
7.8055 | -88.0614
7.8059 | -88.0626
7.8062 | -88.0622
7.8066 | -88.0629
7.8069 | -88.0634
7.8073 | -88.0638
7.8077 | -88.0647
7.808 | -88.0655
7.8084 | -88.0666
7.8088 | -88.0679
7.8091 | -88.0688
7.8095 | -88.0698
7.8098 | -88.0709
7.8102 | -88.0728
7.8106 | -88.0757
7.8109 | -88.0773
7.8113 | -88.0781
7.8116 | -88.0793
7.812 | -88.08
7.8124 | -88.0804
7.8127 | -88.0805
7.8131 | -88.081
7.8135 | -88.0818
7.8138 | -88.0825
7.8142 | -88.0838
7.8145 | -88.0845
7.8149 | -88.0849
7.8153 | -88.0866
7.8156 | -88.0874
7.816 | -88.0881
7.8163 | -88.0884
7.8167 | -88.0895
7.8171 | -88.0905
7.8174 | -88.0913
7.8178 | -88.0918
7.8182 | -88.0914
7.8185 | -88.091
7.8189 | -88.0921
7.8192 | -88.0932
7.8196 | -88.0932
7.82 | -88.0932
7.8203 | -88.0943
7.8207 | -88.0954
7.8211 | -88.096
7.8214 | -88.0967
7.8218 | -88.0976
7.8221 | -88.0988
7.8225 | -88.0997
7.8229 | -88.1003
7.8232 | -88.1012
7.8236 | -88.1018
7.8239 | -88.1024
7.8243 | -88.103
7.8247 | -88.1037
7.825 | -88.1045
7.8254 | -88.105
7.8258 | -88.1056
7.8261 | -88.1065
7.8265 | -88.1072
7.8268 | -88.108
7.8272 | -88.1086
7.8276 | -88.1102
7.8279 | -88.1121
7.8283 | -88.1139
7.8287 | -88.1148
7.829 | -88.115
7.8294 | -88.1155
7.8297 | -88.1161
7.8301 | -88.1172
7.8305 | -88.1182
7.8308 | -88.119
7.8312 | -88.1197
7.8315 | -88.1215
7.8319 | -88.1224
7.8323 | -88.1226
7.8326 | -88.1227
7.833 | -88.1231
7.8334 | -88.1246
7.8337 | -88.1257
7.8341 | -88.1261
7.8344 | -88.1264
7.8348 | -88.127
7.8352 | -88.1277
7.8355 | -88.1281
7.8359 | -88.1291
7.8362 | -88.1305
7.8366 | -88.1306
7.837 | -88.1304
7.8373 | -88.1305
7.8377 | -88.1298
7.8381 | -88.1303
7.8384 | -88.1311
7.8388 | -88.1315
7.8391 | -88.1321
7.8395 | -88.1326
7.8399 | -88.1337
7.8402 | -88.1341
7.8406 | -88.1352
7.841 | -88.1367
7.8413 | -88.1373
7.8417 | -88.1373
7.842 | -88.1374
7.8424 | -88.1376
7.8428 | -88.1377
7.8431 | -88.1378
7.8435 | -88.1382
7.8438 | -88.1386
7.8442 | -88.1392
7.8446 | -88.1393
7.8449 | -88.1398
7.8453 | -88.1399
7.8457 | -88.14
7.846 | -88.1399
7.8464 | -88.1397
7.8467 | -88.1402
7.8471 | -88.1408
7.8475 | -88.1405
7.8478 | -88.1403
7.8482 | -88.1406
7.8486 | -88.1408
7.8489 | -88.1411
7.8493 | -88.1414
7.8496 | -88.1418
7.85 | -88.142
7.8504 | -88.142
7.8507 | -88.142
7.8511 | -88.1424
7.8514 | -88.143
7.8518 | -88.1432
7.8522 | -88.1433
7.8525 | -88.1437
7.8529 | -88.1444
7.8533 | -88.1445
7.8536 | -88.145
7.854 | -88.1456
7.8543 | -88.1455
7.8547 | -88.1456
7.8551 | -88.146
7.8554 | -88.1462
7.8558 | -88.1468
7.8561 | -88.1482
7.8565 | -88.1494
7.8569 | -88.1495
7.8572 | -88.1495
7.8576 | -88.1493
7.858 | -88.149
7.8583 | -88.1486
7.8587 | -88.1488
7.859 | -88.149
7.8594 | -88.1494
7.8598 | -88.1501
7.8601 | -88.15
7.8605 | -88.1503
7.8609 | -88.1508
7.8612 | -88.151
7.8616 | -88.1511
7.8619 | -88.1513
7.8623 | -88.1518
7.8627 | -88.1527
7.863 | -88.153
7.8634 | -88.1527
7.8637 | -88.1528
7.8641 | -88.1532
7.8645 | -88.1539
7.8648 | -88.1539
7.8652 | -88.1542
7.8656 | -88.1547
7.8659 | -88.1549
7.8663 | -88.1553
7.8666 | -88.155
7.867 | -88.155
7.8674 | -88.1554
7.8677 | -88.1559
7.8681 | -88.1577
7.8684 | -88.1582
7.8688 | -88.1582
7.8692 | -88.1596
7.8695 | -88.1609
7.8699 | -88.1604
7.8703 | -88.1604
7.8706 | -88.1603
7.871 | -88.1606
7.8713 | -88.161
7.8717 | -88.1616
7.8721 | -88.1617
7.8724 | -88.1615
7.8728 | -88.1614
7.8732 | -88.161
7.8735 | -88.161
7.8739 | -88.1611
7.8742 | -88.1609
7.8746 | -88.161
7.875 | -88.161
7.8753 | -88.1613
7.8757 | -88.1617
7.876 | -88.1626
7.8764 | -88.1633
7.8768 | -88.1634
7.8771 | -88.1641
7.8775 | -88.1647
7.8779 | -88.1652
7.8782 | -88.1653
7.8786 | -88.1654
7.8789 | -88.1659
7.8793 | -88.1658
7.8797 | -88.1656
7.88 | -88.1657
7.8804 | -88.1658
7.8807 | -88.1666
7.8811 | -88.166
7.8815 | -88.1662
7.8818 | -88.166
7.8822 | -88.1657
7.8826 | -88.1659
7.8829 | -88.1668
7.8833 | -88.1675
7.8836 | -88.1679
7.884 | -88.1679
7.8844 | -88.1676
7.8847 | -88.1673
7.8851 | -88.1682
7.8855 | -88.1694
7.8858 | -88.1698
7.8862 | -88.1699
7.8865 | -88.1699
7.8869 | -88.1701
7.8873 | -88.1706
7.8876 | -88.1708
7.888 | -88.1714
7.8883 | -88.1698
7.8887 | -88.1676
7.8891 | -88.1669
7.8894 | -88.1665
7.8898 | -88.1662
7.8902 | -88.1665
7.8905 | -88.1665
7.8909 | -88.1666
7.8912 | -88.1677
7.8916 | -88.1677
7.892 | -88.168
7.8923 | -88.1685
7.8927 | -88.1689
7.8931 | -88.1686
7.8934 | -88.1687
7.8938 | -88.1684
7.8941 | -88.1681
7.8945 | -88.1678
7.8949 | -88.1676
7.8952 | -88.168
7.8956 | -88.1679
7.8959 | -88.1675
7.8963 | -88.1674
7.8967 | -88.1673
7.897 | -88.167
7.8974 | -88.1668
7.8978 | -88.1662
7.8981 | -88.166
7.8985 | -88.1644
7.8988 | -88.1632
7.8992 | -88.1636
7.8996 | -88.1634
7.8999 | -88.164
7.9003 | -88.1648
7.9006 | -88.1641
7.901 | -88.1626
7.9014 | -88.1612
7.9017 | -88.1595
7.9021 | -88.1585
7.9025 | -88.1581
7.9028 | -88.1583
7.9032 | -88.1584
7.9035 | -88.1573
7.9039 | -88.1565
7.9043 | -88.1563
7.9046 | -88.1561
7.905 | -88.156
7.9054 | -88.1552
7.9057 | -88.1547
7.9061 | -88.1548
7.9064 | -88.155
7.9068 | -88.1548
7.9072 | -88.1553
7.9075 | -88.1552
7.9079 | -88.155
7.9082 | -88.1549
7.9086 | -88.1545
7.909 | -88.1542
7.9093 | -88.1539
7.9097 | -88.1531
7.9101 | -88.1525
7.9104 | -88.1521
7.9108 | -88.1518
7.9111 | -88.1517
7.9115 | -88.1509
7.9119 | -88.1505
7.9122 | -88.15
7.9126 | -88.1487
7.913 | -88.1487
7.9133 | -88.1483
7.9137 | -88.148
7.914 | -88.1482
7.9144 | -88.148
7.9148 | -88.1474
7.9151 | -88.1468
7.9155 | -88.1465
7.9158 | -88.1463
7.9162 | -88.1458
7.9166 | -88.1455
7.9169 | -88.1445
7.9173 | -88.144
7.9177 | -88.1446
7.918 | -88.1451
7.9184 | -88.1449
7.9187 | -88.1444
7.9191 | -88.1438
7.9195 | -88.1433
7.9198 | -88.143
7.9202 | -88.1429
7.9205 | -88.1429
7.9209 | -88.1427
7.9213 | -88.1422
7.9216 | -88.1415
7.922 | -88.1413
7.9224 | -88.1408
7.9227 | -88.1405
7.9231 | -88.1403
7.9234 | -88.1405
7.9238 | -88.1408
7.9242 | -88.1406
7.9245 | -88.1409
7.9249 | -88.1411
7.9253 | -88.1405
7.9256 | -88.1404
7.926 | -88.1403
7.9263 | -88.1402
7.9267 | -88.14
7.9271 | -88.1395
7.9274 | -88.1391
7.9278 | -88.1381
7.9281 | -88.1374
7.9285 | -88.1373
7.9289 | -88.1376
7.9292 | -88.1372
7.9296 | -88.1364
7.93 | -88.1356
7.9303 | -88.1348
7.9307 | -88.1343
7.931 | -88.1338
7.9314 | -88.1333
7.9318 | -88.1329
7.9321 | -88.1322
7.9325 | -88.1315
7.9329 | -88.1321
7.9332 | -88.132
7.9336 | -88.132
7.9339 | -88.1318
7.9343 | -88.1307
7.9347 | -88.13
7.935 | -88.1306
7.9354 | -88.1306
7.9357 | -88.13
7.9361 | -88.1299
7.9365 | -88.1297
7.9368 | -88.1294
7.9372 | -88.1287
7.9376 | -88.128
7.9379 | -88.1273
7.9383 | -88.1267
7.9386 | -88.1259
7.939 | -88.1251
7.9394 | -88.1245
7.9397 | -88.1241
7.9401 | -88.1234
7.9404 | -88.1227
7.9408 | -88.1222
7.9412 | -88.1222
7.9415 | -88.1218
7.9419 | -88.1214
7.9423 | -88.1211
7.9426 | -88.1209
7.943 | -88.1208
7.9433 | -88.1208
7.9437 | -88.1203
7.9441 | -88.1199
7.9444 | -88.1195
7.9448 | -88.1189
7.9452 | -88.1185
7.9455 | -88.1185
7.9459 | -88.118
7.9462 | -88.118
7.9466 | -88.1179
7.947 | -88.1174
7.9473 | -88.117
7.9477 | -88.1167
7.948 | -88.1165
7.9484 | -88.1157
7.9488 | -88.1139
7.9491 | -88.1121
7.9495 | -88.1118
7.9499 | -88.1116
7.9502 | -88.1116
7.9506 | -88.1111
7.9509 | -88.1109
7.9513 | -88.1111
7.9517 | -88.1109
7.952 | -88.1101
7.9524 | -88.1101
7.9527 | -88.1099
7.9531 | -88.109
7.9535 | -88.1077
7.9538 | -88.107
7.9542 | -88.1069
7.9546 | -88.1064
7.9549 | -88.1055
7.9553 | -88.105
7.9556 | -88.1046
7.956 | -88.1041
7.9564 | -88.1038
7.9567 | -88.1037
7.9571 | -88.1033
7.9575 | -88.1029
7.9578 | -88.1024
7.9582 | -88.1018
7.9585 | -88.1012
7.9589 | -88.1008
7.9593 | -88.1005
7.9596 | -88.1002
7.96 | -88.0997
7.9603 | -88.0992
7.9607 | -88.0986
7.9611 | -88.0983
7.9614 | -88.0976
7.9618 | -88.0968
7.9622 | -88.0957
7.9625 | -88.0954
7.9629 | -88.0947
7.9632 | -88.0943
7.9636 | -88.0942
7.964 | -88.0938
7.9643 | -88.0929
7.9647 | -88.0919
7.965 | -88.0915
7.9654 | -88.091
7.9658 | -88.0904
7.9661 | -88.0899
7.9665 | -88.0894
7.9669 | -88.0896
7.9672 | -88.0901
7.9676 | -88.0905
7.9679 | -88.0894
7.9683 | -88.089
7.9687 | -88.0895
7.969 | -88.0899
7.9694 | -88.0894
7.9698 | -88.0886
7.9701 | -88.0882
7.9705 | -88.0885
7.9708 | -88.0882
7.9712 | -88.0877
7.9716 | -88.0868
7.9719 | -88.0855
7.9723 | -88.0846
7.9726 | -88.0835
7.973 | -88.082
7.9734 | -88.0814
7.9737 | -88.0809
7.9741 | -88.0803
7.9745 | -88.0797
7.9748 | -88.0792
7.9752 | -88.0788
7.9755 | -88.0769
7.9759 | -88.0757
7.9763 | -88.0752
7.9766 | -88.0747
7.977 | -88.0747
7.9774 | -88.0747
7.9777 | -88.0742
7.9781 | -88.0742
7.9784 | -88.0738
7.9788 | -88.0736
7.9792 | -88.0731
7.9795 | -88.0728
7.9799 | -88.0721
7.9802 | -88.0714
7.9806 | -88.0706
7.981 | -88.0701
7.9813 | -88.0694
7.9817 | -88.0688
7.9821 | -88.068
7.9824 | -88.0663
7.9828 | -88.0651
7.9831 | -88.0646
7.9835 | -88.0643
7.9839 | -88.0635
7.9842 | -88.0627
7.9846 | -88.0628
7.9849 | -88.0631
7.9853 | -88.0624
7.9857 | -88.0622
7.986 | -88.0613
7.9864 | -88.0607
7.9868 | -88.0602
7.9871 | -88.0601
7.9875 | -88.0598
7.9878 | -88.0592
7.9882 | -88.0594
7.9886 | -88.0591
7.9889 | -88.0589
7.9893 | -88.0591
7.9897 | -88.0586
7.99 | -88.0591
7.9904 | -88.06
7.9907 | -88.0598
7.9911 | -88.0593
7.9915 | -88.0591
7.9918 | -88.0586
7.9922 | -88.0582
7.9925 | -88.0578
7.9929 | -88.0574
7.9933 | -88.0568
7.9936 | -88.057
7.994 | -88.0568
7.9944 | -88.0561
7.9947 | -88.0557
7.9951 | -88.0555
7.9954 | -88.0553
7.9958 | -88.055
7.9962 | -88.0547
7.9965 | -88.0549
7.9969 | -88.0544
7.9973 | -88.0532
7.9976 | -88.0527
7.998 | -88.0525
7.9983 | -88.0519
7.9987 | -88.0511
7.9991 | -88.0508
7.9994 | -88.0506
7.9998 | -88.0504
8.0001 | -88.0506
8.0005 | -88.0504
8.0009 | -88.0498
8.0012 | -88.0496
8.0016 | -88.0493
8.002 | -88.0497
8.0023 | -88.0507
8.0027 | -88.0513
8.003 | -88.052
8.0034 | -88.052
8.0038 | -88.0517
8.0041 | -88.0518
8.0045 | -88.0514
8.0048 | -88.0513
8.0052 | -88.0512
8.0056 | -88.0508
8.0059 | -88.0503
8.0063 | -88.0501
8.0067 | -88.0487
8.007 | -88.0478
8.0074 | -88.0472
8.0077 | -88.0466
8.0081 | -88.0463
8.0085 | -88.0457
8.0088 | -88.0452
8.0092 | -88.0442
8.0096 | -88.0435
8.0099 | -88.0434
8.0103 | -88.0431
8.0106 | -88.0428
8.011 | -88.0426
8.0114 | -88.0428
8.0117 | -88.0434
8.0121 | -88.0437
8.0124 | -88.0439
8.0128 | -88.0439
8.0132 | -88.0433
8.0135 | -88.0424
8.0139 | -88.0421
8.0143 | -88.0418
8.0146 | -88.0415
8.015 | -88.0417
8.0153 | -88.0419
8.0157 | -88.0417
8.0161 | -88.0412
8.0164 | -88.0409
8.0168 | -88.0408
8.0171 | -88.0407
8.0175 | -88.0412
8.0179 | -88.041
8.0182 | -88.0407
8.0186 | -88.0406
8.019 | -88.0412
8.0193 | -88.042
8.0197 | -88.042
8.02 | -88.0424
8.0204 | -88.0422
8.0208 | -88.0426
8.0211 | -88.0431
8.0215 | -88.0425
8.0219 | -88.0423
8.0222 | -88.0423
8.0226 | -88.0424
8.0229 | -88.042
8.0233 | -88.0412
8.0237 | -88.0408
8.024 | -88.0408
8.0244 | -88.0405
8.0247 | -88.0405
8.0251 | -88.0404
8.0255 | -88.0403
8.0258 | -88.0407
8.0262 | -88.0408
8.0266 | -88.0406
8.0269 | -88.0402
8.0273 | -88.0408
8.0276 | -88.0403
8.028 | -88.0408
8.0284 | -88.041
8.0287 | -88.0408
8.0291 | -88.0412
8.0295 | -88.0413
8.0298 | -88.0411
8.0302 | -88.0409
8.0305 | -88.0408
8.0309 | -88.0409
8.0313 | -88.041
8.0316 | -88.0405
8.032 | -88.0397
8.0323 | -88.0392
8.0327 | -88.0391
8.0331 | -88.0394
8.0334 | -88.0399
8.0338 | -88.0399
8.0342 | -88.0396
8.0345 | -88.0395
8.0349 | -88.0393
8.0352 | -88.0387
8.0356 | -88.0382
8.036 | -88.0388
8.0363 | -88.0393
8.0367 | -88.0395
8.037 | -88.0397
8.0374 | -88.0393
8.0378 | -88.0394
8.0381 | -88.0389
8.0385 | -88.039
8.0389 | -88.0393
8.0392 | -88.0393
8.0396 | -88.0391
8.0399 | -88.039
8.0403 | -88.0387
8.0407 | -88.0387
8.041 | -88.0392
8.0414 | -88.0398
8.0418 | -88.04
8.0421 | -88.0401
8.0425 | -88.0397
8.0428 | -88.0391
8.0432 | -88.0394
8.0436 | -88.0398
8.0439 | -88.0402
8.0443 | -88.0404
8.0446 | -88.0405
8.045 | -88.0415
8.0454 | -88.0424
8.0457 | -88.0429
8.0461 | -88.0432
8.0465 | -88.0437
8.0468 | -88.0441
8.0472 | -88.0438
8.0475 | -88.0437
8.0479 | -88.0439
8.0483 | -88.0448
8.0486 | -88.0448
8.049 | -88.0446
8.0494 | -88.0448
8.0497 | -88.0447
8.0501 | -88.0446
8.0504 | -88.0447
8.0508 | -88.0448
8.0512 | -88.0454
8.0515 | -88.046
8.0519 | -88.0464
8.0522 | -88.0472
8.0526 | -88.0473
8.053 | -88.047
8.0533 | -88.0472
8.0537 | -88.0472
8.0541 | -88.0477
8.0544 | -88.0476
8.0548 | -88.0476
8.0551 | -88.0488
8.0555 | -88.0509
8.0559 | -88.0514
8.0562 | -88.0515
8.0566 | -88.0514
8.0569 | -88.0516
8.0573 | -88.0521
8.0577 | -88.0521
8.058 | -88.0523
8.0584 | -88.0527
8.0588 | -88.0538
8.0591 | -88.054
8.0595 | -88.0547
8.0598 | -88.0553
8.0602 | -88.0553
8.0606 | -88.0556
8.0609 | -88.056
8.0613 | -88.0563
8.0617 | -88.0567
8.062 | -88.0567
8.0624 | -88.0568
8.0627 | -88.0565
8.0631 | -88.0564
8.0635 | -88.0573
8.0638 | -88.0574
8.0642 | -88.0575
8.0645 | -88.058
8.0649 | -88.059
8.0653 | -88.0598
8.0656 | -88.06
8.066 | -88.0606
8.0664 | -88.0609
8.0667 | -88.0615
8.0671 | -88.062
8.0674 | -88.0626
8.0678 | -88.0634
8.0682 | -88.0635
8.0685 | -88.0636
8.0689 | -88.065
8.0692 | -88.0659
8.0696 | -88.0675
8.07 | -88.0679
8.0703 | -88.0684
8.0707 | -88.0689
8.0711 | -88.0693
8.0714 | -88.0695
8.0718 | -88.0698
8.0721 | -88.0701
8.0725 | -88.0703
8.0729 | -88.0708
8.0732 | -88.0712
8.0736 | -88.0709
8.074 | -88.0716
8.0743 | -88.0724
8.0747 | -88.0732
8.075 | -88.0737
8.0754 | -88.074
8.0758 | -88.0747
8.0761 | -88.0752
8.0765 | -88.0756
8.0768 | -88.0764
8.0772 | -88.0772
8.0776 | -88.0777
8.0779 | -88.0783
8.0783 | -88.0789
8.0787 | -88.0794
8.079 | -88.0799
8.0794 | -88.0815
8.0797 | -88.0819
8.0801 | -88.0822
8.0805 | -88.0829
8.0808 | -88.0838
8.0812 | -88.0845
8.0816 | -88.0844
8.0819 | -88.0841
8.0823 | -88.0845
8.0826 | -88.0847
8.083 | -88.0838
8.0834 | -88.0833
8.0837 | -88.0837
8.0841 | -88.0845
8.0844 | -88.0852
8.0848 | -88.086
8.0852 | -88.087
8.0855 | -88.0874
8.0859 | -88.0884
8.0863 | -88.0896
8.0866 | -88.0902
8.087 | -88.0908
8.0873 | -88.0913
8.0877 | -88.092
8.0881 | -88.0925
8.0884 | -88.0932
8.0888 | -88.0939
8.0891 | -88.0945
8.0895 | -88.0951
8.0899 | -88.0968
8.0902 | -88.0963
8.0906 | -88.0962
8.091 | -88.0968
8.0913 | -88.0975
8.0917 | -88.098
8.092 | -88.0985
8.0924 | -88.0986
8.0928 | -88.0994
8.0931 | -88.0998
8.0935 | -88.1002
8.0939 | -88.1008
8.0942 | -88.1017
8.0946 | -88.1028
8.0949 | -88.1037
8.0953 | -88.1043
8.0957 | -88.1048
8.096 | -88.1056
8.0964 | -88.1064
8.0967 | -88.1077
8.0971 | -88.1087
8.0975 | -88.1094
8.0978 | -88.1099
8.0982 | -88.1102
8.0986 | -88.1112
8.0989 | -88.1122
8.0993 | -88.1129
8.0996 | -88.1139
8.1 | -88.1151
8.1004 | -88.1168
8.1007 | -88.1176
8.1011 | -88.1181
8.1014 | -88.1191
8.1018 | -88.1198
8.1022 | -88.121
8.1025 | -88.1225
8.1029 | -88.123
8.1033 | -88.1233
8.1036 | -88.1232
8.104 | -88.1234
8.1043 | -88.124
8.1047 | -88.1245
8.1051 | -88.1254
8.1054 | -88.1258
8.1058 | -88.1262
8.1062 | -88.1272
8.1065 | -88.128
8.1069 | -88.1291
8.1072 | -88.1302
8.1076 | -88.1311
8.108 | -88.1315
8.1083 | -88.132
8.1087 | -88.1328
8.109 | -88.1337
8.1094 | -88.1346
8.1098 | -88.1354
8.1101 | -88.1362
8.1105 | -88.1369
8.1109 | -88.1387
8.1112 | -88.1395
8.1116 | -88.1404
8.1119 | -88.1411
8.1123 | -88.1418
8.1127 | -88.1423
8.113 | -88.1433
8.1134 | -88.1443
8.1137 | -88.1453
8.1141 | -88.1461
8.1145 | -88.1468
8.1148 | -88.1475
8.1152 | -88.1487
8.1156 | -88.1496
8.1159 | -88.1504
8.1163 | -88.151
8.1166 | -88.1521
8.117 | -88.1527
8.1174 | -88.1537
8.1177 | -88.1551
8.1181 | -88.1559
8.1185 | -88.1573
8.1188 | -88.1583
8.1192 | -88.1589
8.1195 | -88.1598
8.1199 | -88.1615
8.1203 | -88.1634
8.1206 | -88.1651
8.121 | -88.1662
8.1213 | -88.1678
8.1217 | -88.1687
8.1221 | -88.17
8.1224 | -88.1711
8.1228 | -88.1717
8.1232 | -88.1731
8.1235 | -88.1741
8.1239 | -88.1745
8.1242 | -88.1755
8.1246 | -88.1761
8.125 | -88.1771
8.1253 | -88.1781
8.1257 | -88.1793
8.1261 | -88.1804
8.1264 | -88.1813
8.1268 | -88.1825
8.1271 | -88.1833
8.1275 | -88.1837
8.1279 | -88.1847
8.1282 | -88.1854
8.1286 | -88.1854
8.1289 | -88.1861
8.1293 | -88.1867
8.1297 | -88.1869
8.13 | -88.1886
8.1304 | -88.1907
8.1308 | -88.1935
8.1311 | -88.1964
8.1315 | -88.1989
8.1318 | -88.2007
8.1322 | -88.2025
8.1326 | -88.2034
8.1329 | -88.2034
8.1333 | -88.2044
8.1336 | -88.205
8.134 | -88.2048
8.1344 | -88.2054
8.1347 | -88.2064
8.1351 | -88.207
8.1355 | -88.2077
8.1358 | -88.2075
8.1362 | -88.2079
8.1365 | -88.2094
8.1369 | -88.2096
8.1373 | -88.2078
8.1376 | -88.2064
8.138 | -88.2063
8.1384 | -88.2072
8.1387 | -88.2084
8.1391 | -88.2105
8.1394 | -88.2119
8.1398 | -88.2133
8.1402 | -88.2147
8.1405 | -88.216
8.1409 | -88.2168
8.1412 | -88.2178
8.1416 | -88.2182
8.142 | -88.2187
8.1423 | -88.2187
8.1427 | -88.2196
8.1431 | -88.2206
8.1434 | -88.2219
8.1438 | -88.2226
8.1441 | -88.2228
8.1445 | -88.2229
8.1449 | -88.2237
8.1452 | -88.2252
8.1456 | -88.2264
8.146 | -88.2278
8.1463 | -88.2291
8.1467 | -88.23
8.147 | -88.2308
8.1474 | -88.2315
8.1478 | -88.2326
8.1481 | -88.2331
8.1485 | -88.2339
8.1488 | -88.235
8.1492 | -88.2361
8.1496 | -88.2371
8.1499 | -88.2384
8.1503 | -88.2397
8.1507 | -88.2406
8.151 | -88.2413
8.1514 | -88.2423
8.1517 | -88.2436
8.1521 | -88.2443
8.1525 | -88.2458
8.1528 | -88.2468
8.1532 | -88.2477
8.1535 | -88.2487
8.1539 | -88.2496
8.1543 | -88.2502
8.1546 | -88.2509
8.155 | -88.2519
8.1554 | -88.2529
8.1557 | -88.2536
8.1561 | -88.2543
8.1564 | -88.2555
8.1568 | -88.2567
8.1572 | -88.2577
8.1575 | -88.2587
8.1579 | -88.2598
8.1583 | -88.2613
8.1586 | -88.2622
8.159 | -88.2628
8.1593 | -88.2637
8.1597 | -88.2647
8.1601 | -88.2658
8.1604 | -88.2671
8.1608 | -88.2686
8.1611 | -88.2705
8.1615 | -88.2725
8.1619 | -88.2739
8.1622 | -88.275
8.1626 | -88.2762
8.163 | -88.2776
8.1633 | -88.2789
8.1637 | -88.2803
8.164 | -88.2816
8.1644 | -88.2826
8.1648 | -88.2838
8.1651 | -88.2854
8.1655 | -88.2867
8.1659 | -88.2878
8.1662 | -88.2891
8.1666 | -88.29
8.1669 | -88.2916
8.1673 | -88.2921
8.1677 | -88.2924
8.168 | -88.2938
8.1684 | -88.2949
8.1687 | -88.296
8.1691 | -88.2972
8.1695 | -88.2985
8.1698 | -88.2999
8.1702 | -88.3006
8.1706 | -88.3014
8.1709 | -88.3025
8.1713 | -88.3043
8.1716 | -88.3061
8.172 | -88.308
8.1724 | -88.3096
8.1727 | -88.3101
8.1731 | -88.3108
8.1734 | -88.3119
8.1738 | -88.3129
8.1742 | -88.3142
8.1745 | -88.3156
8.1749 | -88.317
8.1753 | -88.3181
8.1756 | -88.3194
8.176 | -88.321
8.1763 | -88.3219
8.1767 | -88.3228
8.1771 | -88.3243
8.1774 | -88.3256
8.1778 | -88.3271
8.1782 | -88.3284
8.1785 | -88.3297
8.1789 | -88.3305
8.1792 | -88.3311
8.1796 | -88.3325
8.18 | -88.3336
8.1803 | -88.3342
8.1807 | -88.3352
8.181 | -88.3358
8.1814 | -88.3362
8.1818 | -88.337
8.1821 | -88.3382
8.1825 | -88.3395
8.1829 | -88.3398
8.1832 | -88.3402
8.1836 | -88.3414
8.1839 | -88.3424
8.1843 | -88.3434
8.1847 | -88.3449
8.185 | -88.3462
8.1854 | -88.3471
8.1857 | -88.3485
8.1861 | -88.3499
8.1865 | -88.3511
8.1868 | -88.3523
8.1872 | -88.3529
8.1876 | -88.3534
8.1879 | -88.3546
8.1883 | -88.3558
8.1886 | -88.3568
8.189 | -88.3572
8.1894 | -88.3578
8.1897 | -88.3591
8.1901 | -88.36
8.1905 | -88.3615
8.1908 | -88.363
8.1912 | -88.3643
8.1915 | -88.3657
8.1919 | -88.367
8.1923 | -88.3688
8.1926 | -88.3708
8.193 | -88.3722
8.1933 | -88.3733
8.1937 | -88.3746
8.1941 | -88.3757
8.1944 | -88.377
8.1948 | -88.3781
8.1952 | -88.3793
8.1955 | -88.3806
8.1959 | -88.3821
8.1962 | -88.3839
8.1966 | -88.3851
8.197 | -88.3863
8.1973 | -88.387
8.1977 | -88.3878
8.1981 | -88.3891
8.1984 | -88.3895
8.1988 | -88.389
8.1991 | -88.3898
8.1995 | -88.3909
8.1999 | -88.3921
8.2002 | -88.3926
8.2006 | -88.3935
8.2009 | -88.3945
8.2013 | -88.3955
8.2017 | -88.3963
8.202 | -88.3972
8.2024 | -88.3989
8.2028 | -88.4006
8.2031 | -88.4014
8.2035 | -88.4022
8.2038 | -88.4035
8.2042 | -88.4048
8.2046 | -88.4059
8.2049 | -88.4069
8.2053 | -88.4083
8.2056 | -88.41
8.206 | -88.4112
8.2064 | -88.4122
8.2067 | -88.4137
8.2071 | -88.4148
8.2075 | -88.4171
8.2078 | -88.4192
8.2082 | -88.4206
8.2085 | -88.4215
8.2089 | -88.4225
8.2093 | -88.4236
8.2096 | -88.425
8.21 | -88.4268
8.2104 | -88.4279
8.2107 | -88.4289
8.2111 | -88.4297
8.2114 | -88.4303
8.2118 | -88.4313
8.2122 | -88.4325
8.2125 | -88.434
8.2129 | -88.4353
8.2132 | -88.4365
8.2136 | -88.4379
8.214 | -88.4401
8.2143 | -88.4412
8.2147 | -88.4419
8.2151 | -88.442
8.2154 | -88.4429
8.2158 | -88.444
8.2161 | -88.4447
8.2165 | -88.4458
8.2169 | -88.4467
8.2172 | -88.4476
8.2176 | -88.4495
8.2179 | -88.4512
8.2183 | -88.4521
8.2187 | -88.4532
8.219 | -88.4543
8.2194 | -88.4554
8.2198 | -88.4566
8.2201 | -88.4571
8.2205 | -88.458
8.2208 | -88.4585
8.2212 | -88.4589
8.2216 | -88.4594
8.2219 | -88.4601
8.2223 | -88.4611
8.2227 | -88.4622
8.223 | -88.4636
8.2234 | -88.4648
8.2237 | -88.4659
8.2241 | -88.4668
8.2245 | -88.4677
8.2248 | -88.4684
8.2252 | -88.4692
8.2255 | -88.4702
8.2259 | -88.4709
8.2263 | -88.4718
8.2266 | -88.4728
8.227 | -88.4738
8.2274 | -88.4746
8.2277 | -88.4757
8.2281 | -88.4776
8.2284 | -88.4795
8.2288 | -88.4804
8.2292 | -88.4811
8.2295 | -88.4819
8.2299 | -88.4832
8.2302 | -88.4845
8.2306 | -88.4857
8.231 | -88.4879
8.2313 | -88.4894
8.2317 | -88.4898
8.2321 | -88.4902
8.2324 | -88.4902
8.2328 | -88.4908
8.2331 | -88.4915
8.2335 | -88.4925
8.2339 | -88.4934
8.2342 | -88.4942
8.2346 | -88.4951
8.235 | -88.4964
8.2353 | -88.4975
8.2357 | -88.4983
8.236 | -88.4984
8.2364 | -88.4983
8.2368 | -88.4989
8.2371 | -88.4995
8.2375 | -88.4999
8.2378 | -88.5003
8.2382 | -88.5012
8.2386 | -88.5023
8.2389 | -88.503
8.2393 | -88.5045
8.2397 | -88.5057
8.24 | -88.5066
8.2404 | -88.5078
8.2407 | -88.5087
8.2411 | -88.5095
8.2415 | -88.5103
8.2418 | -88.5111
8.2422 | -88.5122
8.2426 | -88.514
8.2429 | -88.515
8.2433 | -88.5159
8.2436 | -88.5168
8.244 | -88.5174
8.2444 | -88.5182
8.2447 | -88.5194
8.2451 | -88.5203
8.2454 | -88.5213
8.2458 | -88.5221
8.2462 | -88.5234
8.2465 | -88.524
8.2469 | -88.5251
8.2473 | -88.5261
8.2476 | -88.5272
8.248 | -88.5282
8.2483 | -88.5292
8.2487 | -88.5299
8.2491 | -88.5308
8.2494 | -88.5318
8.2498 | -88.5327
8.2501 | -88.5334
8.2505 | -88.5341
8.2509 | -88.5346
8.2512 | -88.5356
8.2516 | -88.5377
8.252 | -88.5404
8.2523 | -88.541
8.2527 | -88.5419
8.253 | -88.5432
8.2534 | -88.5439
8.2538 | -88.5447
8.2541 | -88.5454
8.2545 | -88.5462
8.2549 | -88.5468
8.2552 | -88.5472
8.2556 | -88.5478
8.2559 | -88.5484
8.2563 | -88.5491
8.2567 | -88.5498
8.257 | -88.5504
8.2574 | -88.5513
8.2577 | -88.5527
8.2581 | -88.5541
8.2585 | -88.5549
8.2588 | -88.5558
8.2592 | -88.5564
8.2596 | -88.5572
8.2599 | -88.5578
8.2603 | -88.5584
8.2606 | -88.5591
8.261 | -88.5604
8.2614 | -88.5613
8.2617 | -88.5619
8.2621 | -88.5625
8.2624 | -88.5636
8.2628 | -88.5652
8.2632 | -88.5665
8.2635 | -88.5671
8.2639 | -88.5674
8.2643 | -88.5682
8.2646 | -88.5689
8.265 | -88.5692
8.2653 | -88.57
8.2657 | -88.5713
8.2661 | -88.5722
8.2664 | -88.5729
8.2668 | -88.5735
8.2672 | -88.5746
8.2675 | -88.5754
8.2679 | -88.5772
8.2682 | -88.5777
8.2686 | -88.5776
8.269 | -88.5777
8.2693 | -88.5784
8.2697 | -88.579
8.27 | -88.5797
8.2704 | -88.5809
8.2708 | -88.5819
8.2711 | -88.5824
8.2715 | -88.5829
8.2719 | -88.5832
8.2722 | -88.5839
8.2726 | -88.5848
8.2729 | -88.5853
8.2733 | -88.5854
8.2737 | -88.5857
8.274 | -88.5856
8.2744 | -88.5855
8.2748 | -88.5853
8.2751 | -88.5854
8.2755 | -88.586
8.2758 | -88.5871
8.2762 | -88.5877
8.2766 | -88.5883
8.2769 | -88.5889
8.2773 | -88.5894
8.2776 | -88.5898
8.278 | -88.5901
8.2784 | -88.5911
8.2787 | -88.5932
8.2791 | -88.5943
8.2795 | -88.595
8.2798 | -88.5961
8.2802 | -88.5971
8.2805 | -88.5973
8.2809 | -88.5978
8.2813 | -88.5986
8.2816 | -88.5991
8.282 | -88.5997
8.2823 | -88.5998
8.2827 | -88.5992
8.2831 | -88.6003
8.2834 | -88.6013
8.2838 | -88.6019
8.2842 | -88.6025
8.2845 | -88.6029
8.2849 | -88.6037
8.2852 | -88.6043
8.2856 | -88.6052
8.286 | -88.6056
8.2863 | -88.6056
8.2867 | -88.6054
8.2871 | -88.6056
8.2874 | -88.6059
8.2878 | -88.6061
8.2881 | -88.6057
8.2885 | -88.6059
8.2889 | -88.6063
8.2892 | -88.606
8.2896 | -88.6063
8.2899 | -88.6068
8.2903 | -88.6071
8.2907 | -88.6075
8.291 | -88.6088
8.2914 | -88.6096
8.2918 | -88.6091
8.2921 | -88.6091
8.2925 | -88.6101
8.2928 | -88.6111
8.2932 | -88.6118
8.2936 | -88.6122
8.2939 | -88.6126
8.2943 | -88.6134
8.2947 | -88.6138
8.295 | -88.6142
8.2954 | -88.6146
8.2957 | -88.615
8.2961 | -88.6152
8.2965 | -88.6157
8.2968 | -88.6162
8.2972 | -88.6166
8.2975 | -88.6181
8.2979 | -88.6188
8.2983 | -88.6183
8.2986 | -88.617
8.299 | -88.6176
8.2994 | -88.6182
8.2997 | -88.6187
8.3001 | -88.6191
8.3004 | -88.6196
8.3008 | -88.6201
8.3012 | -88.6204
8.3015 | -88.6208
8.3019 | -88.6212
8.3022 | -88.6223
8.3026 | -88.6231
8.303 | -88.623
8.3033 | -88.6228
8.3037 | -88.6234
8.3041 | -88.624
8.3044 | -88.6256
8.3048 | -88.6268
8.3051 | -88.6275
8.3055 | -88.6278
8.3059 | -88.6281
8.3062 | -88.6281
8.3066 | -88.6282
8.307 | -88.6284
8.3073 | -88.6287
8.3077 | -88.6287
8.308 | -88.629
8.3084 | -88.6294
8.3088 | -88.6306
8.3091 | -88.6311
8.3095 | -88.6318
8.3098 | -88.6323
8.3102 | -88.6334
8.3106 | -88.6338
8.3109 | -88.6345
8.3113 | -88.6354
8.3117 | -88.636
8.312 | -88.6364
8.3124 | -88.6365
8.3127 | -88.6366
8.3131 | -88.6369
8.3135 | -88.6369
8.3138 | -88.6365
8.3142 | -88.6358
8.3146 | -88.6358
8.3149 | -88.636
8.3153 | -88.6364
8.3156 | -88.6368
8.316 | -88.6374
8.3164 | -88.6374
8.3167 | -88.6382
8.3171 | -88.6384
8.3174 | -88.638
8.3178 | -88.6382
8.3182 | -88.6383
8.3185 | -88.6388
8.3189 | -88.6397
8.3193 | -88.64
8.3196 | -88.6378
8.32 | -88.6378
8.3203 | -88.6376
8.3207 | -88.6375
8.3211 | -88.6395
8.3214 | -88.6404
8.3218 | -88.6406
8.3221 | -88.6411
8.3225 | -88.6415
8.3229 | -88.6429
8.3232 | -88.6435
8.3236 | -88.6438
8.324 | -88.6445
8.3243 | -88.6455
8.3247 | -88.6473
8.325 | -88.6484
8.3254 | -88.6483
8.3258 | -88.6484
8.3261 | -88.6478
8.3265 | -88.6469
8.3269 | -88.6462
8.3272 | -88.6467
8.3276 | -88.6479
8.3279 | -88.6474
8.3283 | -88.6484
8.3287 | -88.6503
8.329 | -88.6507
8.3294 | -88.6514
8.3297 | -88.6526
8.3301 | -88.6522
8.3305 | -88.6533
8.3308 | -88.6541
8.3312 | -88.654
8.3316 | -88.6538
8.3319 | -88.6539
8.3323 | -88.6537
8.3326 | -88.6531
8.333 | -88.653
8.3334 | -88.6529
8.3337 | -88.6529
8.3341 | -88.6522
8.3344 | -88.651
8.3348 | -88.6505
8.3352 | -88.651
8.3355 | -88.6529
8.3359 | -88.6543
8.3363 | -88.6545
8.3366 | -88.6542
8.337 | -88.6541
8.3373 | -88.654
8.3377 | -88.6548
8.3381 | -88.6557
8.3384 | -88.6556
8.3388 | -88.6564
8.3392 | -88.6576
8.3395 | -88.6577
8.3399 | -88.6574
8.3402 | -88.6577
8.3406 | -88.6579
8.341 | -88.6575
8.3413 | -88.6567
8.3417 | -88.6566
8.342 | -88.6574
8.3424 | -88.6579
8.3428 | -88.658
8.3431 | -88.6579
8.3435 | -88.6579
8.3439 | -88.6576
8.3442 | -88.6573
8.3446 | -88.6569
8.3449 | -88.6567
8.3453 | -88.6575
8.3457 | -88.6578
8.346 | -88.6576
8.3464 | -88.6577
8.3468 | -88.6575
8.3471 | -88.6572
8.3475 | -88.6572
8.3478 | -88.6572
8.3482 | -88.6572
8.3486 | -88.6562
8.3489 | -88.6546
8.3493 | -88.6538
8.3496 | -88.6532
8.35 | -88.6527
8.3504 | -88.6523
8.3507 | -88.6505
8.3511 | -88.65
8.3515 | -88.6494
8.3518 | -88.6488
8.3522 | -88.6483
8.3525 | -88.648
8.3529 | -88.648
8.3533 | -88.6479
8.3536 | -88.6479
8.354 | -88.6477
8.3543 | -88.6476
8.3547 | -88.6475
8.3551 | -88.6475
8.3554 | -88.6475
8.3558 | -88.6475
8.3562 | -88.6469
8.3565 | -88.6456
8.3569 | -88.645
8.3572 | -88.6443
8.3576 | -88.644
8.358 | -88.6446
8.3583 | -88.6443
8.3587 | -88.6441
8.3591 | -88.6448
8.3594 | -88.6444
8.3598 | -88.6447
8.3601 | -88.6442
8.3605 | -88.6442
8.3609 | -88.6449
8.3612 | -88.6446
8.3616 | -88.6449
8.3619 | -88.6448
8.3623 | -88.6444
8.3627 | -88.6442
8.363 | -88.6441
8.3634 | -88.6439
8.3638 | -88.644
8.3641 | -88.6443
8.3645 | -88.6443
8.3648 | -88.6439
8.3652 | -88.6444
8.3656 | -88.6453
8.3659 | -88.6462
8.3663 | -88.6457
8.3666 | -88.6445
8.367 | -88.6442
8.3674 | -88.6446
8.3677 | -88.644
8.3681 | -88.6443
8.3685 | -88.6449
8.3688 | -88.6441
8.3692 | -88.6434
8.3695 | -88.643
8.3699 | -88.6421
8.3703 | -88.6411
8.3706 | -88.6404
8.371 | -88.64
8.3714 | -88.6391
8.3717 | -88.6391
8.3721 | -88.6392
8.3724 | -88.6396
8.3728 | -88.6395
8.3732 | -88.6391
8.3735 | -88.6383
8.3739 | -88.6381
8.3742 | -88.6381
8.3746 | -88.6385
8.375 | -88.6386
8.3753 | -88.6373
8.3757 | -88.6363
8.3761 | -88.636
8.3764 | -88.6358
8.3768 | -88.636
8.3771 | -88.6356
8.3775 | -88.6354
8.3779 | -88.6355
8.3782 | -88.6357
8.3786 | -88.6354
8.3789 | -88.6348
8.3793 | -88.6344
8.3797 | -88.6347
8.38 | -88.6357
8.3804 | -88.6364
8.3808 | -88.636
8.3811 | -88.6354
8.3815 | -88.6351
8.3818 | -88.6354
8.3822 | -88.6355
8.3826 | -88.6362
8.3829 | -88.636
8.3833 | -88.6353
8.3837 | -88.6347
8.384 | -88.6343
8.3844 | -88.6338
8.3847 | -88.6333
8.3851 | -88.6333
8.3855 | -88.6335
8.3858 | -88.6328
8.3862 | -88.6316
8.3865 | -88.6308
8.3869 | -88.6305
8.3873 | -88.6303
8.3876 | -88.6305
8.388 | -88.6308
8.3884 | -88.6308
8.3887 | -88.6305
8.3891 | -88.6304
8.3894 | -88.6298
8.3898 | -88.6298
8.3902 | -88.6292
8.3905 | -88.6284
8.3909 | -88.6278
8.3913 | -88.6273
8.3916 | -88.6272
8.392 | -88.6272
8.3923 | -88.627
8.3927 | -88.6264
8.3931 | -88.6258
8.3934 | -88.6242
8.3938 | -88.6224
8.3941 | -88.6219
8.3945 | -88.6214
8.3949 | -88.6202
8.3952 | -88.6191
8.3956 | -88.619
8.396 | -88.6191
8.3963 | -88.6188
8.3967 | -88.6185
8.397 | -88.6181
8.3974 | -88.6175
8.3978 | -88.6174
8.3981 | -88.6171
8.3985 | -88.6159
8.3988 | -88.615
8.3992 | -88.6144
8.3996 | -88.614
8.3999 | -88.6142
8.4003 | -88.6147
8.4007 | -88.6153
8.401 | -88.616
8.4014 | -88.6163
8.4017 | -88.6161
8.4021 | -88.6157
8.4025 | -88.6145
8.4028 | -88.6133
8.4032 | -88.6126
8.4036 | -88.6118
8.4039 | -88.6113
8.4043 | -88.6114
8.4046 | -88.6115
8.405 | -88.611
8.4054 | -88.6108
8.4057 | -88.611
8.4061 | -88.6111
8.4064 | -88.6112
8.4068 | -88.6111
8.4072 | -88.6112
8.4075 | -88.6105
8.4079 | -88.6097
8.4083 | -88.6094
8.4086 | -88.6091
8.409 | -88.6087
8.4093 | -88.6088
8.4097 | -88.6086
8.4101 | -88.609
8.4104 | -88.6094
8.4108 | -88.6094
8.4112 | -88.6094
8.4115 | -88.6097
8.4119 | -88.6099
8.4122 | -88.6091
8.4126 | -88.6085
8.413 | -88.6082
8.4133 | -88.608
8.4137 | -88.6073
8.414 | -88.6069
8.4144 | -88.6066
8.4148 | -88.6065
8.4151 | -88.6064
8.4155 | -88.6067
8.4159 | -88.6068
8.4162 | -88.6064
8.4166 | -88.606
8.4169 | -88.6058
8.4173 | -88.6055
8.4177 | -88.6052
8.418 | -88.6046
8.4184 | -88.6034
8.4187 | -88.6018
8.4191 | -88.6008
8.4195 | -88.6005
8.4198 | -88.6
8.4202 | -88.5993
8.4206 | -88.599
8.4209 | -88.5986
8.4213 | -88.5979
8.4216 | -88.5975
8.422 | -88.5976
8.4224 | -88.5978
8.4227 | -88.5982
8.4231 | -88.5979
8.4235 | -88.5972
8.4238 | -88.597
8.4242 | -88.5973
8.4245 | -88.5968
8.4249 | -88.5965
8.4253 | -88.5964
8.4256 | -88.5961
8.426 | -88.5962
8.4263 | -88.5957
8.4267 | -88.5956
8.4271 | -88.5953
8.4274 | -88.5953
8.4278 | -88.595
8.4282 | -88.5944
8.4285 | -88.5936
8.4289 | -88.5933
8.4292 | -88.5928
8.4296 | -88.5924
8.43 | -88.5923
8.4303 | -88.5917
8.4307 | -88.5908
8.4311 | -88.5904
8.4314 | -88.59
8.4318 | -88.5899
8.4321 | -88.5899
8.4325 | -88.5897
8.4329 | -88.5893
8.4332 | -88.5886
8.4336 | -88.5874
8.4339 | -88.5868
8.4343 | -88.5866
8.4347 | -88.5859
8.435 | -88.5851
8.4354 | -88.5847
8.4358 | -88.5849
8.4361 | -88.5846
8.4365 | -88.5844
8.4368 | -88.5838
8.4372 | -88.5834
8.4376 | -88.583
8.4379 | -88.5822
8.4383 | -88.5814
8.4386 | -88.5812
8.439 | -88.5812
8.4394 | -88.5816
8.4397 | -88.5816
8.4401 | -88.5811
8.4405 | -88.581
8.4408 | -88.5804
8.4412 | -88.58
8.4415 | -88.5789
8.4419 | -88.5783
8.4423 | -88.5784
8.4426 | -88.5788
8.443 | -88.5788
8.4434 | -88.5786
8.4437 | -88.5781
8.4441 | -88.5776
8.4444 | -88.5772
8.4448 | -88.5764
8.4452 | -88.5761
8.4455 | -88.5756
8.4459 | -88.5752
8.4462 | -88.5749
8.4466 | -88.574
8.447 | -88.5735
8.4473 | -88.5731
8.4477 | -88.5721
8.4481 | -88.5715
8.4484 | -88.5716
8.4488 | -88.5716
8.4491 | -88.5715
8.4495 | -88.5716
8.4499 | -88.5715
8.4502 | -88.571
8.4506 | -88.5705
8.4509 | -88.5698
8.4513 | -88.5691
8.4517 | -88.5687
8.452 | -88.568
8.4524 | -88.5671
8.4528 | -88.5656
8.4531 | -88.5639
8.4535 | -88.5635
8.4538 | -88.5633
8.4542 | -88.5629
8.4546 | -88.5624
8.4549 | -88.5626
8.4553 | -88.5623
8.4557 | -88.5636
8.456 | -88.5655
8.4564 | -88.5652
8.4567 | -88.563
8.4571 | -88.562
8.4575 | -88.5616
8.4578 | -88.5605
8.4582 | -88.5599
8.4585 | -88.5595
8.4589 | -88.559
8.4593 | -88.5584
8.4596 | -88.5578
8.46 | -88.5574
8.4604 | -88.5572
8.4607 | -88.5575
8.4611 | -88.5574
8.4614 | -88.5573
8.4618 | -88.5572
8.4622 | -88.5563
8.4625 | -88.5562
8.4629 | -88.5561
8.4633 | -88.556
8.4636 | -88.5558
8.464 | -88.5554
8.4643 | -88.5549
8.4647 | -88.5545
8.4651 | -88.5545
8.4654 | -88.5544
8.4658 | -88.5543
8.4661 | -88.5538
8.4665 | -88.5551
8.4669 | -88.5561
8.4672 | -88.5556
8.4676 | -88.5555
8.468 | -88.5554
8.4683 | -88.5548
8.4687 | -88.5552
8.469 | -88.5558
8.4694 | -88.5565
8.4698 | -88.5574
8.4701 | -88.5575
8.4705 | -88.557
8.4708 | -88.5565
8.4712 | -88.5561
8.4716 | -88.5556
8.4719 | -88.5552
8.4723 | -88.555
8.4727 | -88.5552
8.473 | -88.5548
8.4734 | -88.554
8.4737 | -88.5536
8.4741 | -88.553
8.4745 | -88.5524
8.4748 | -88.552
8.4752 | -88.5515
8.4756 | -88.5511
8.4759 | -88.5506
8.4763 | -88.5494
8.4766 | -88.549
8.477 | -88.5485
8.4774 | -88.5482
8.4777 | -88.5476
8.4781 | -88.5472
8.4784 | -88.5474
8.4788 | -88.5467
8.4792 | -88.5459
8.4795 | -88.5457
8.4799 | -88.5447
8.4803 | -88.5435
8.4806 | -88.5424
8.481 | -88.542
8.4813 | -88.5422
8.4817 | -88.542
8.4821 | -88.5415
8.4824 | -88.5407
8.4828 | -88.5396
8.4831 | -88.5386
8.4835 | -88.5381
8.4839 | -88.5376
8.4842 | -88.5371
8.4846 | -88.5367
8.485 | -88.536
8.4853 | -88.5356
8.4857 | -88.5351
8.486 | -88.5354
8.4864 | -88.5348
8.4868 | -88.5341
8.4871 | -88.5339
8.4875 | -88.5333
8.4879 | -88.5325
8.4882 | -88.5319
8.4886 | -88.5317
8.4889 | -88.5305
8.4893 | -88.5296
8.4897 | -88.5287
8.49 | -88.5281
8.4904 | -88.5276
8.4907 | -88.527
8.4911 | -88.5265
8.4915 | -88.5255
8.4918 | -88.5239
8.4922 | -88.5227
8.4926 | -88.522
8.4929 | -88.5215
8.4933 | -88.5208
8.4936 | -88.5203
8.494 | -88.5203
8.4944 | -88.52
8.4947 | -88.5196
8.4951 | -88.5192
8.4954 | -88.5188
8.4958 | -88.5184
8.4962 | -88.5181
8.4965 | -88.5178
8.4969 | -88.5175
8.4973 | -88.5168
8.4976 | -88.5163
8.498 | -88.5167
8.4983 | -88.5169
8.4987 | -88.5165
8.4991 | -88.516
8.4994 | -88.5154
8.4998 | -88.5153
8.5002 | -88.5154
8.5005 | -88.5153
8.5009 | -88.5148
8.5012 | -88.5143
8.5016 | -88.514
8.502 | -88.5135
8.5023 | -88.5127
8.5027 | -88.5118
8.503 | -88.5111
8.5034 | -88.5106
8.5038 | -88.5101
8.5041 | -88.5096
8.5045 | -88.5094
8.5049 | -88.5091
8.5052 | -88.5091
8.5056 | -88.5088
8.5059 | -88.5083
8.5063 | -88.5074
8.5067 | -88.5067
8.507 | -88.5061
8.5074 | -88.5049
8.5078 | -88.5046
8.5081 | -88.5041
8.5085 | -88.5039
8.5088 | -88.5037
8.5092 | -88.5032
8.5096 | -88.5029
8.5099 | -88.5024
8.5103 | -88.5019
8.5106 | -88.5012
8.511 | -88.5002
8.5114 | -88.4993
8.5117 | -88.4985
8.5121 | -88.4975
8.5125 | -88.4969
8.5128 | -88.4964
8.5132 | -88.4959
8.5135 | -88.4952
8.5139 | -88.4947
8.5143 | -88.494
8.5146 | -88.4933
8.515 | -88.4928
8.5153 | -88.4923
8.5157 | -88.4918
8.5161 | -88.4914
8.5164 | -88.4911
8.5168 | -88.4908
8.5172 | -88.4904
8.5175 | -88.4901
8.5179 | -88.4893
8.5182 | -88.4886
8.5186 | -88.4879
8.519 | -88.4871
8.5193 | -88.4866
8.5197 | -88.4859
8.5201 | -88.4855
8.5204 | -88.4865
8.5208 | -88.4877
8.5211 | -88.4872
8.5215 | -88.4872
8.5219 | -88.4871
8.5222 | -88.4859
8.5226 | -88.4845
8.5229 | -88.484
8.5233 | -88.4837
8.5237 | -88.4835
8.524 | -88.4828
8.5244 | -88.4817
8.5248 | -88.4815
8.5251 | -88.4821
8.5255 | -88.4819
8.5258 | -88.481
8.5262 | -88.4805
8.5266 | -88.4804
8.5269 | -88.4806
8.5273 | -88.48
8.5277 | -88.4795
8.528 | -88.4789
8.5284 | -88.4784
8.5287 | -88.4779
8.5291 | -88.4774
8.5295 | -88.4769
8.5298 | -88.4765
8.5302 | -88.4761
8.5305 | -88.4752
8.5309 | -88.4746
8.5313 | -88.4744
8.5316 | -88.4741
8.532 | -88.4732
8.5324 | -88.4722
8.5327 | -88.4714
8.5331 | -88.471
8.5334 | -88.4704
8.5338 | -88.4701
8.5342 | -88.4699
8.5345 | -88.4702
8.5349 | -88.47
8.5352 | -88.4691
8.5356 | -88.4682
8.536 | -88.4678
8.5363 | -88.4665
8.5367 | -88.4657
8.5371 | -88.4656
8.5374 | -88.4652
8.5378 | -88.4644
8.5381 | -88.4629
8.5385 | -88.4622
8.5389 | -88.4608
8.5392 | -88.4586
8.5396 | -88.457
8.54 | -88.4555
8.5403 | -88.455
8.5407 | -88.4547
8.541 | -88.4549
8.5414 | -88.454
8.5418 | -88.4522
8.5421 | -88.4504
8.5425 | -88.4489
8.5428 | -88.4476
8.5432 | -88.4459
8.5436 | -88.4446
8.5439 | -88.4439
8.5443 | -88.4435
8.5447 | -88.4431
8.545 | -88.4429
8.5454 | -88.442
8.5457 | -88.442
8.5461 | -88.4423
8.5465 | -88.4419
8.5468 | -88.4413
8.5472 | -88.4407
8.5475 | -88.4401
8.5479 | -88.4393
8.5483 | -88.4387
8.5486 | -88.4386
8.549 | -88.4395
8.5494 | -88.4393
8.5497 | -88.4391
8.5501 | -88.4386
8.5504 | -88.4376
8.5508 | -88.436
8.5512 | -88.4348
8.5515 | -88.434
8.5519 | -88.4331
8.5523 | -88.4321
8.5526 | -88.4315
8.553 | -88.4311
8.5533 | -88.4304
8.5537 | -88.4289
8.5541 | -88.4283
8.5544 | -88.428
8.5548 | -88.4275
8.5551 | -88.4274
8.5555 | -88.4269
8.5559 | -88.4264
8.5562 | -88.4269
8.5566 | -88.4269
8.557 | -88.4262
8.5573 | -88.4251
8.5577 | -88.4249
8.558 | -88.4246
8.5584 | -88.4239
8.5588 | -88.423
8.5591 | -88.4224
8.5595 | -88.4213
8.5599 | -88.4179
8.5602 | -88.4164
8.5606 | -88.416
8.5609 | -88.4155
8.5613 | -88.4148
8.5617 | -88.4141
8.562 | -88.4137
8.5624 | -88.4126
8.5627 | -88.411
8.5631 | -88.4099
8.5635 | -88.4094
8.5638 | -88.4087
8.5642 | -88.4082
8.5646 | -88.4077
8.5649 | -88.4067
8.5653 | -88.4058
8.5656 | -88.4044
8.566 | -88.4034
8.5664 | -88.4027
8.5667 | -88.4019
8.5671 | -88.401
8.5674 | -88.4003
8.5678 | -88.3999
8.5682 | -88.3991
8.5685 | -88.3984
8.5689 | -88.398
8.5693 | -88.3973
8.5696 | -88.3962
8.57 | -88.3957
8.5703 | -88.3948
8.5707 | -88.3942
8.5711 | -88.3938
8.5714 | -88.3928
8.5718 | -88.392
8.5722 | -88.3915
8.5725 | -88.3915
8.5729 | -88.3911
8.5732 | -88.3904
8.5736 | -88.39
8.574 | -88.3894
8.5743 | -88.3887
8.5747 | -88.3881
8.575 | -88.3875
8.5754 | -88.3866
8.5758 | -88.3861
8.5761 | -88.3856
8.5765 | -88.3849
8.5769 | -88.3837
8.5772 | -88.3824
8.5776 | -88.3813
8.5779 | -88.3803
8.5783 | -88.3793
8.5787 | -88.3785
8.579 | -88.3776
8.5794 | -88.3766
8.5798 | -88.3757
8.5801 | -88.3751
8.5805 | -88.3744
8.5808 | -88.3737
8.5812 | -88.3732
8.5816 | -88.3724
8.5819 | -88.3712
8.5823 | -88.3708
8.5826 | -88.3713
8.583 | -88.3715
8.5834 | -88.3713
8.5837 | -88.3709
8.5841 | -88.3706
8.5845 | -88.3699
8.5848 | -88.3685
8.5852 | -88.3681
8.5855 | -88.3677
8.5859 | -88.367
8.5863 | -88.3664
8.5866 | -88.3658
8.587 | -88.3649
8.5873 | -88.3639
8.5877 | -88.3631
8.5881 | -88.3621
8.5884 | -88.361
8.5888 | -88.3607
8.5892 | -88.3609
8.5895 | -88.3606
8.5899 | -88.3597
8.5902 | -88.3584
8.5906 | -88.3574
8.591 | -88.3567
8.5913 | -88.3559
8.5917 | -88.3552
8.5921 | -88.3541
8.5924 | -88.3522
8.5928 | -88.3502
8.5931 | -88.349
8.5935 | -88.3484
8.5939 | -88.3477
8.5942 | -88.347
8.5946 | -88.3462
8.5949 | -88.3455
8.5953 | -88.3447
8.5957 | -88.3441
8.596 | -88.3438
8.5964 | -88.3437
8.5968 | -88.3431
8.5971 | -88.3425
8.5975 | -88.3417
8.5978 | -88.3408
8.5982 | -88.3401
8.5986 | -88.3397
8.5989 | -88.3393
8.5993 | -88.3385
8.5996 | -88.3382
8.6 | -88.3382
8.6004 | -88.3377
8.6007 | -88.3371
8.6011 | -88.3364
8.6015 | -88.3348
8.6018 | -88.334
8.6022 | -88.3334
8.6025 | -88.3329
8.6029 | -88.3325
8.6033 | -88.3323
8.6036 | -88.3316
8.604 | -88.3312
8.6044 | -88.3308
8.6047 | -88.3302
8.6051 | -88.3298
8.6054 | -88.3289
8.6058 | -88.3278
8.6062 | -88.3273
8.6065 | -88.3267
8.6069 | -88.3258
8.6072 | -88.325
8.6076 | -88.3244
8.608 | -88.3238
8.6083 | -88.3233
8.6087 | -88.3229
8.6091 | -88.3228
8.6094 | -88.3228
8.6098 | -88.3224
8.6101 | -88.3219
8.6105 | -88.3214
8.6109 | -88.321
8.6112 | -88.3205
8.6116 | -88.3199
8.612 | -88.3195
8.6123 | -88.3189
8.6127 | -88.3184
8.613 | -88.3184
8.6134 | -88.3181
8.6138 | -88.3174
8.6141 | -88.3172
8.6145 | -88.3165
8.6148 | -88.3158
8.6152 | -88.3154
8.6156 | -88.315
8.6159 | -88.3148
8.6163 | -88.3146
8.6167 | -88.3139
8.617 | -88.3134
8.6174 | -88.3128
8.6177 | -88.3122
8.6181 | -88.3115
8.6185 | -88.3112
8.6188 | -88.3106
8.6192 | -88.3101
8.6195 | -88.3096
8.6199 | -88.3088
8.6203 | -88.3078
8.6206 | -88.3074
8.621 | -88.3076
8.6214 | -88.3068
8.6217 | -88.3057
8.6221 | -88.3051
8.6224 | -88.305
8.6228 | -88.3042
8.6232 | -88.3034
8.6235 | -88.3027
8.6239 | -88.3021
8.6243 | -88.3017
8.6246 | -88.3012
8.625 | -88.3007
8.6253 | -88.3
8.6257 | -88.2997
8.6261 | -88.2994
8.6264 | -88.2989
8.6268 | -88.298
8.6271 | -88.2971
8.6275 | -88.2967
8.6279 | -88.2964
8.6282 | -88.2955
8.6286 | -88.2942
8.629 | -88.2936
8.6293 | -88.2929
8.6297 | -88.2928
8.63 | -88.2922
8.6304 | -88.2922
8.6308 | -88.2925
8.6311 | -88.2918
8.6315 | -88.2909
8.6318 | -88.2904
8.6322 | -88.29
8.6326 | -88.2901
8.6329 | -88.2895
8.6333 | -88.2891
8.6337 | -88.2888
8.634 | -88.2884
8.6344 | -88.2881
8.6347 | -88.2875
8.6351 | -88.2867
8.6355 | -88.2861
8.6358 | -88.286
8.6362 | -88.2859
8.6366 | -88.2852
8.6369 | -88.2844
8.6373 | -88.2836
8.6376 | -88.2831
8.638 | -88.2827
8.6384 | -88.2825
8.6387 | -88.2819
8.6391 | -88.2815
8.6394 | -88.2811
8.6398 | -88.2807
8.6402 | -88.2803
8.6405 | -88.2799
8.6409 | -88.2794
8.6413 | -88.2788
8.6416 | -88.2784
8.642 | -88.2782
8.6423 | -88.2782
8.6427 | -88.2786
8.6431 | -88.2785
8.6434 | -88.2781
8.6438 | -88.2777
8.6442 | -88.2772
8.6445 | -88.2771
8.6449 | -88.277
8.6452 | -88.2765
8.6456 | -88.2757
8.646 | -88.2748
8.6463 | -88.2743
8.6467 | -88.2738
8.647 | -88.273
8.6474 | -88.2724
8.6478 | -88.2722
8.6481 | -88.272
8.6485 | -88.2718
8.6489 | -88.2714
8.6492 | -88.271
8.6496 | -88.2706
8.6499 | -88.2705
8.6503 | -88.2706
8.6507 | -88.2702
8.651 | -88.2697
8.6514 | -88.2693
8.6517 | -88.2692
8.6521 | -88.2687
8.6525 | -88.2681
8.6528 | -88.267
8.6532 | -88.2665
8.6536 | -88.2667
8.6539 | -88.266
8.6543 | -88.2663
8.6546 | -88.2663
8.655 | -88.2663
8.6554 | -88.2654
8.6557 | -88.2644
8.6561 | -88.2641
8.6565 | -88.2637
8.6568 | -88.2633
8.6572 | -88.2627
8.6575 | -88.2621
8.6579 | -88.2618
8.6583 | -88.2615
8.6586 | -88.2612
8.659 | -88.2609
8.6593 | -88.2605
8.6597 | -88.2597
8.6601 | -88.2594
8.6604 | -88.2592
8.6608 | -88.259
8.6612 | -88.2591
8.6615 | -88.2589
8.6619 | -88.2586
8.6622 | -88.2584
8.6626 | -88.2587
8.663 | -88.2592
8.6633 | -88.2586
8.6637 | -88.2583
8.664 | -88.258
8.6644 | -88.259
8.6648 | -88.2597
8.6651 | -88.2596
8.6655 | -88.2593
8.6659 | -88.2589
8.6662 | -88.2582
8.6666 | -88.2579
8.6669 | -88.2578
8.6673 | -88.2577
8.6677 | -88.2575
8.668 | -88.2574
8.6684 | -88.2573
8.6688 | -88.2574
8.6691 | -88.2574
8.6695 | -88.2574
8.6698 | -88.2569
8.6702 | -88.2562
8.6706 | -88.2557
8.6709 | -88.2553
8.6713 | -88.2551
8.6716 | -88.255
8.672 | -88.2558
8.6724 | -88.2566
8.6727 | -88.2567
8.6731 | -88.2566
8.6735 | -88.2566
8.6738 | -88.2567
8.6742 | -88.2569
8.6745 | -88.257
8.6749 | -88.2564
8.6753 | -88.2565
8.6756 | -88.2562
8.676 | -88.2562
8.6764 | -88.2558
8.6767 | -88.2556
8.6771 | -88.2554
8.6774 | -88.2553
8.6778 | -88.2553
8.6782 | -88.2555
8.6785 | -88.2556
8.6789 | -88.256
8.6792 | -88.2562
8.6796 | -88.256
8.68 | -88.256
8.6803 | -88.2563
8.6807 | -88.2563
8.6811 | -88.2557
8.6814 | -88.2551
8.6818 | -88.2551
8.6821 | -88.255
8.6825 | -88.2549
8.6829 | -88.255
8.6832 | -88.255
8.6836 | -88.2551
8.6839 | -88.2551
8.6843 | -88.2552
8.6847 | -88.255
8.685 | -88.2553
8.6854 | -88.2548
8.6858 | -88.2551
8.6861 | -88.2555
8.6865 | -88.2563
8.6868 | -88.2568
8.6872 | -88.2567
8.6876 | -88.2566
8.6879 | -88.2564
8.6883 | -88.257
8.6887 | -88.2572
8.689 | -88.2571
8.6894 | -88.2566
8.6897 | -88.2558
8.6901 | -88.2551
8.6905 | -88.2546
8.6908 | -88.2545
8.6912 | -88.2549
8.6915 | -88.2543
8.6919 | -88.2539
8.6923 | -88.2541
8.6926 | -88.2544
8.693 | -88.2547
8.6934 | -88.2549
8.6937 | -88.2547
8.6941 | -88.2547
8.6944 | -88.2547
8.6948 | -88.2544
8.6952 | -88.2543
8.6955 | -88.2547
8.6959 | -88.2547
8.6962 | -88.2551
8.6966 | -88.255
8.697 | -88.2551
8.6973 | -88.2551
8.6977 | -88.255
8.6981 | -88.2549
8.6984 | -88.255
8.6988 | -88.2549
8.6991 | -88.2549
8.6995 | -88.2552
8.6999 | -88.2556
8.7002 | -88.2555
8.7006 | -88.2556
8.701 | -88.2558
8.7013 | -88.2556
8.7017 | -88.2558
8.702 | -88.2558
8.7024 | -88.256
8.7028 | -88.2565
8.7031 | -88.2567
8.7035 | -88.257
8.7038 | -88.2574
8.7042 | -88.2573
8.7046 | -88.2576
8.7049 | -88.2581
8.7053 | -88.2581
8.7057 | -88.2591
8.706 | -88.2592
8.7064 | -88.2589
8.7067 | -88.259
8.7071 | -88.2594
8.7075 | -88.2592
8.7078 | -88.2593
8.7082 | -88.2599
8.7086 | -88.2599
8.7089 | -88.2606
8.7093 | -88.2608
8.7096 | -88.261
8.71 | -88.2612
8.7104 | -88.2612
8.7107 | -88.2604
8.7111 | -88.2604
8.7114 | -88.2608
8.7118 | -88.2617
8.7122 | -88.2616
8.7125 | -88.262
8.7129 | -88.2632
8.7133 | -88.264
8.7136 | -88.2649
8.714 | -88.2661
8.7143 | -88.2665
8.7147 | -88.2667
8.7151 | -88.2669
8.7154 | -88.2673
8.7158 | -88.2675
8.7161 | -88.2686
8.7165 | -88.2687
8.7169 | -88.269
8.7172 | -88.2699
8.7176 | -88.2707
8.718 | -88.2712
8.7183 | -88.2707
8.7187 | -88.2705
8.719 | -88.271
8.7194 | -88.2702
8.7198 | -88.2696
8.7201 | -88.2701
8.7205 | -88.271
8.7209 | -88.2729
8.7212 | -88.2739
8.7216 | -88.2741
8.7219 | -88.2751
8.7223 | -88.278
8.7227 | -88.2806
8.723 | -88.2818
8.7234 | -88.2821
8.7237 | -88.2824
8.7241 | -88.2828
8.7245 | -88.2833
8.7248 | -88.283
8.7252 | -88.2819
8.7256 | -88.2815
8.7259 | -88.2818
8.7263 | -88.2822
8.7266 | -88.2834
8.727 | -88.2842
8.7274 | -88.2848
8.7277 | -88.2861
8.7281 | -88.2871
8.7285 | -88.2871
8.7288 | -88.2873
8.7292 | -88.2875
8.7295 | -88.2883
8.7299 | -88.2891
8.7303 | -88.2902
8.7306 | -88.2913
8.731 | -88.292
8.7313 | -88.2922
8.7317 | -88.2918
8.7321 | -88.292
8.7324 | -88.2921
8.7328 | -88.2924
8.7332 | -88.2928
8.7335 | -88.2929
8.7339 | -88.2938
8.7342 | -88.2946
8.7346 | -88.2952
8.735 | -88.2955
8.7353 | -88.2956
8.7357 | -88.2953
8.736 | -88.2955
8.7364 | -88.2962
8.7368 | -88.297
8.7371 | -88.2978
8.7375 | -88.2985
8.7379 | -88.2988
8.7382 | -88.2994
8.7386 | -88.2998
8.7389 | -88.2994
8.7393 | -88.2994
8.7397 | -88.2997
8.74 | -88.3001
8.7404 | -88.3007
8.7408 | -88.3016
8.7411 | -88.3021
8.7415 | -88.3021
8.7418 | -88.3022
8.7422 | -88.3013
8.7426 | -88.3011
8.7429 | -88.3014
8.7433 | -88.3017
8.7436 | -88.3026
8.744 | -88.3031
8.7444 | -88.3034
8.7447 | -88.3038
8.7451 | -88.3048
8.7455 | -88.3058
8.7458 | -88.3061
8.7462 | -88.3066
8.7465 | -88.3069
8.7469 | -88.3072
8.7473 | -88.3068
8.7476 | -88.3073
8.748 | -88.3076
8.7483 | -88.308
8.7487 | -88.3078
8.7491 | -88.3078
8.7494 | -88.3079
8.7498 | -88.3082
8.7502 | -88.3088
8.7505 | -88.3093
8.7509 | -88.3094
8.7512 | -88.3097
8.7516 | -88.3103
8.752 | -88.3108
8.7523 | -88.3112
8.7527 | -88.3118
8.7531 | -88.3124
8.7534 | -88.3128
8.7538 | -88.3136
8.7541 | -88.314
8.7545 | -88.3152
8.7549 | -88.3158
8.7552 | -88.3159
8.7556 | -88.3164
8.7559 | -88.3169
8.7563 | -88.3174
8.7567 | -88.3178
8.757 | -88.3185
8.7574 | -88.3198
8.7578 | -88.3206
8.7581 | -88.3216
8.7585 | -88.3222
8.7588 | -88.3228
8.7592 | -88.3224
8.7596 | -88.3224
8.7599 | -88.3226
8.7603 | -88.3237
8.7607 | -88.3243
8.761 | -88.3244
8.7614 | -88.3252
8.7617 | -88.3257
8.7621 | -88.3258
8.7625 | -88.3263
8.7628 | -88.3269
8.7632 | -88.3267
8.7635 | -88.3263
8.7639 | -88.3268
8.7643 | -88.3278
8.7646 | -88.3288
8.765 | -88.3293
8.7654 | -88.3296
8.7657 | -88.3301
8.7661 | -88.3308
8.7664 | -88.3315
8.7668 | -88.3326
8.7672 | -88.3337
8.7675 | -88.3344
8.7679 | -88.335
8.7682 | -88.3361
8.7686 | -88.3374
8.769 | -88.3381
8.7693 | -88.3391
8.7697 | -88.3404
8.7701 | -88.3415
8.7704 | -88.3424
8.7708 | -88.343
8.7711 | -88.3447
8.7715 | -88.3459
8.7719 | -88.3463
8.7722 | -88.3472
8.7726 | -88.3481
8.773 | -88.3485
8.7733 | -88.3487
8.7737 | -88.3492
8.774 | -88.3499
8.7744 | -88.3505
8.7748 | -88.3511
8.7751 | -88.3517
8.7755 | -88.3525
8.7758 | -88.354
8.7762 | -88.3555
8.7766 | -88.3564
8.7769 | -88.3571
8.7773 | -88.3577
8.7777 | -88.3589
8.778 | -88.36
8.7784 | -88.3603
8.7787 | -88.3609
8.7791 | -88.3617
8.7795 | -88.3623
8.7798 | -88.3631
8.7802 | -88.3638
8.7805 | -88.3643
8.7809 | -88.3649
8.7813 | -88.3657
8.7816 | -88.3664
8.782 | -88.3671
8.7824 | -88.368
8.7827 | -88.3688
8.7831 | -88.3699
8.7834 | -88.3706
8.7838 | -88.3716
8.7842 | -88.3721
8.7845 | -88.3724
8.7849 | -88.3733
8.7853 | -88.3748
8.7856 | -88.3759
8.786 | -88.3763
8.7863 | -88.3763
8.7867 | -88.3773
8.7871 | -88.3784
8.7874 | -88.3794
8.7878 | -88.3807
8.7881 | -88.383
8.7885 | -88.3841
8.7889 | -88.3852
8.7892 | -88.3865
8.7896 | -88.3879
8.79 | -88.389
8.7903 | -88.3902
8.7907 | -88.3912
8.791 | -88.392
8.7914 | -88.3927
8.7918 | -88.3939
8.7921 | -88.395
8.7925 | -88.3961
8.7929 | -88.3974
8.7932 | -88.3988
8.7936 | -88.3997
8.7939 | -88.4005
8.7943 | -88.4017
8.7947 | -88.4027
8.795 | -88.4037
8.7954 | -88.4046
8.7957 | -88.4055
8.7961 | -88.4066
8.7965 | -88.4077
8.7968 | -88.4085
8.7972 | -88.4093
8.7976 | -88.4105
8.7979 | -88.4114
8.7983 | -88.4124
8.7986 | -88.4136
8.799 | -88.4146
8.7994 | -88.4156
8.7997 | -88.417
8.8001 | -88.4179
8.8004 | -88.4189
8.8008 | -88.42
8.8012 | -88.4212
8.8015 | -88.4225
8.8019 | -88.4239
8.8023 | -88.425
8.8026 | -88.426
8.803 | -88.4272
8.8033 | -88.4284
8.8037 | -88.4294
8.8041 | -88.4298
8.8044 | -88.4297
8.8048 | -88.4303
8.8052 | -88.4317
8.8055 | -88.4333
8.8059 | -88.4345
8.8062 | -88.4356
8.8066 | -88.4366
8.807 | -88.4378
8.8073 | -88.439
8.8077 | -88.4401
8.808 | -88.4411
8.8084 | -88.4423
8.8088 | -88.4434
8.8091 | -88.4446
8.8095 | -88.4456
8.8099 | -88.4467
8.8102 | -88.4486
8.8106 | -88.4495
8.8109 | -88.4507
8.8113 | -88.4522
8.8117 | -88.4536
8.812 | -88.4551
8.8124 | -88.4565
8.8127 | -88.458
8.8131 | -88.4594
8.8135 | -88.4605
8.8138 | -88.4614
8.8142 | -88.4625
8.8146 | -88.4633
8.8149 | -88.464
8.8153 | -88.4647
8.8156 | -88.4657
8.816 | -88.4671
8.8164 | -88.4683
8.8167 | -88.4694
8.8171 | -88.4706
8.8175 | -88.4715
8.8178 | -88.4724
8.8182 | -88.4734
8.8185 | -88.4738
8.8189 | -88.4747
8.8193 | -88.4759
8.8196 | -88.4764
8.82 | -88.4768
8.8203 | -88.478
8.8207 | -88.4798
8.8211 | -88.4811
8.8214 | -88.482
8.8218 | -88.4834
8.8222 | -88.485
8.8225 | -88.4863
8.8229 | -88.4877
8.8232 | -88.4895
8.8236 | -88.4915
8.824 | -88.4929
8.8243 | -88.4948
8.8247 | -88.4962
8.8251 | -88.4976
8.8254 | -88.4987
8.8258 | -88.5003
8.8261 | -88.5017
8.8265 | -88.503
8.8269 | -88.5044
8.8272 | -88.5059
8.8276 | -88.5075
8.8279 | -88.5085
8.8283 | -88.5094
8.8287 | -88.5107
8.829 | -88.5122
8.8294 | -88.5137
8.8298 | -88.5152
8.8301 | -88.517
8.8305 | -88.5182
8.8308 | -88.5195
8.8312 | -88.5216
8.8316 | -88.523
8.8319 | -88.5242
8.8323 | -88.5255
8.8326 | -88.5264
8.833 | -88.5276
8.8334 | -88.5288
8.8337 | -88.5301
8.8341 | -88.5312
8.8345 | -88.5324
8.8348 | -88.534
8.8352 | -88.5356
8.8355 | -88.5369
8.8359 | -88.5383
8.8363 | -88.5394
8.8366 | -88.5409
8.837 | -88.5422
8.8374 | -88.5434
8.8377 | -88.545
8.8381 | -88.5463
8.8384 | -88.5477
8.8388 | -88.5495
8.8392 | -88.5511
8.8395 | -88.5523
8.8399 | -88.5537
8.8402 | -88.555
8.8406 | -88.5563
8.841 | -88.5577
8.8413 | -88.5588
8.8417 | -88.5597
8.8421 | -88.5611
8.8424 | -88.5623
8.8428 | -88.5637
8.8431 | -88.565
8.8435 | -88.5664
8.8439 | -88.568
8.8442 | -88.5695
8.8446 | -88.5706
8.845 | -88.572
8.8453 | -88.5732
8.8457 | -88.5754
8.846 | -88.5774
8.8464 | -88.5788
8.8468 | -88.58
8.8471 | -88.5814
8.8475 | -88.5826
8.8478 | -88.5838
8.8482 | -88.585
8.8486 | -88.5868
8.8489 | -88.5883
8.8493 | -88.5896
8.8497 | -88.591
8.85 | -88.5924
8.8504 | -88.5938
8.8507 | -88.5949
8.8511 | -88.596
8.8515 | -88.5978
8.8518 | -88.5994
8.8522 | -88.6009
8.8525 | -88.6023
8.8529 | -88.603
8.8533 | -88.6043
8.8536 | -88.6058
8.854 | -88.6072
8.8544 | -88.6088
8.8547 | -88.61
8.8551 | -88.6113
8.8554 | -88.6128
8.8558 | -88.6144
8.8562 | -88.6156
8.8565 | -88.6169
8.8569 | -88.6187
8.8573 | -88.6202
8.8576 | -88.6216
8.858 | -88.6231
8.8583 | -88.6241
8.8587 | -88.6242
8.8591 | -88.6251
8.8594 | -88.6267
8.8598 | -88.6281
8.8601 | -88.6294
8.8605 | -88.6304
8.8609 | -88.632
8.8612 | -88.6335
8.8616 | -88.6349
8.862 | -88.6359
8.8623 | -88.6377
8.8627 | -88.6393
8.863 | -88.6411
8.8634 | -88.6429
8.8638 | -88.6447
8.8641 | -88.6463
8.8645 | -88.6479
8.8648 | -88.6495
8.8652 | -88.6511
8.8656 | -88.6528
8.8659 | -88.6541
8.8663 | -88.6552
8.8667 | -88.6571
8.867 | -88.6584
8.8674 | -88.6598
8.8677 | -88.6615
8.8681 | -88.6636
8.8685 | -88.6656
8.8688 | -88.6678
8.8692 | -88.6701
8.8696 | -88.6725
8.8699 | -88.6744
8.8703 | -88.6761
8.8706 | -88.6775
8.871 | -88.6788
8.8714 | -88.6802
8.8717 | -88.6817
8.8721 | -88.6834
8.8724 | -88.685
8.8728 | -88.6871
8.8732 | -88.6886
8.8735 | -88.6902
8.8739 | -88.6919
8.8743 | -88.6936
8.8746 | -88.6951
8.875 | -88.6968
8.8753 | -88.6986
8.8757 | -88.6998
8.8761 | -88.7008
8.8764 | -88.7025
8.8768 | -88.7039
8.8772 | -88.7049
8.8775 | -88.7066
8.8779 | -88.7082
8.8782 | -88.7103
8.8786 | -88.7121
8.879 | -88.7136
8.8793 | -88.7151
8.8797 | -88.7165
8.88 | -88.7179
8.8804 | -88.7194
8.8808 | -88.7208
8.8811 | -88.7224
8.8815 | -88.7243
8.8819 | -88.7265
8.8822 | -88.7288
8.8826 | -88.7308
8.8829 | -88.7328
8.8833 | -88.7347
8.8837 | -88.7364
8.884 | -88.7381
8.8844 | -88.7402
8.8847 | -88.7416
8.8851 | -88.7432
8.8855 | -88.7455
8.8858 | -88.7471
8.8862 | -88.7486
8.8866 | -88.7503
8.8869 | -88.7522
8.8873 | -88.7542
8.8876 | -88.7561
8.888 | -88.7585
8.8884 | -88.7605
8.8887 | -88.7622
8.8891 | -88.7636
8.8895 | -88.765
8.8898 | -88.7665
8.8902 | -88.7679
8.8905 | -88.769
8.8909 | -88.7705
8.8913 | -88.7724
8.8916 | -88.7747
8.892 | -88.7769
8.8923 | -88.7789
8.8927 | -88.7805
8.8931 | -88.7821
8.8934 | -88.7836
8.8938 | -88.7857
8.8942 | -88.7875
8.8945 | -88.7896
8.8949 | -88.7922
8.8952 | -88.7953
8.8956 | -88.7972
8.896 | -88.7989
8.8963 | -88.8007
8.8967 | -88.8023
8.8971 | -88.8037
8.8974 | -88.8053
8.8978 | -88.8062
8.8981 | -88.8078
8.8985 | -88.8095
8.8989 | -88.811
8.8992 | -88.8125
8.8996 | -88.8141
8.8999 | -88.8161
8.9003 | -88.818
8.9007 | -88.8196
8.901 | -88.8212
8.9014 | -88.8228
8.9018 | -88.8249
8.9021 | -88.8264
8.9025 | -88.8281
8.9028 | -88.8298
8.9032 | -88.8317
8.9036 | -88.8334
8.9039 | -88.8351
8.9043 | -88.8374
8.9046 | -88.8397
8.905 | -88.8419
8.9054 | -88.8434
8.9057 | -88.8451
8.9061 | -88.8469
8.9065 | -88.8483
8.9068 | -88.85
8.9072 | -88.8518
8.9075 | -88.8536
8.9079 | -88.8552
8.9083 | -88.8568
8.9086 | -88.8584
8.909 | -88.8601
8.9094 | -88.8616
8.9097 | -88.863
8.9101 | -88.8648
8.9104 | -88.8664
8.9108 | -88.8683
8.9112 | -88.8704
8.9115 | -88.8728
8.9119 | -88.8757
8.9122 | -88.8776
8.9126 | -88.8789
8.913 | -88.8804
8.9133 | -88.882
8.9137 | -88.8838
8.9141 | -88.8855
8.9144 | -88.8873
8.9148 | -88.8895
8.9151 | -88.8916
8.9155 | -88.8938
8.9159 | -88.8952
8.9162 | -88.8969
8.9166 | -88.8988
8.9169 | -88.9008
8.9173 | -88.9033
8.9177 | -88.9064
8.918 | -88.9082
8.9184 | -88.91
8.9188 | -88.9126
8.9191 | -88.9149
8.9195 | -88.9166
8.9198 | -88.9185
8.9202 | -88.9204
8.9206 | -88.9222
8.9209 | -88.924
8.9213 | -88.9257
8.9217 | -88.928
8.922 | -88.93
8.9224 | -88.9318
8.9227 | -88.9335
8.9231 | -88.9352
8.9235 | -88.9367
8.9238 | -88.9386
8.9242 | -88.94
8.9245 | -88.9422
8.9249 | -88.9441
8.9253 | -88.947
8.9256 | -88.9494
8.926 | -88.9516
8.9264 | -88.9535
8.9267 | -88.9559
8.9271 | -88.9582
8.9274 | -88.96
8.9278 | -88.962
8.9282 | -88.9641
8.9285 | -88.966
8.9289 | -88.9683
8.9292 | -88.9702
8.9296 | -88.971
8.93 | -88.971
8.9303 | -88.9722
8.9307 | -88.9742
8.9311 | -88.9755
8.9314 | -88.9771
8.9318 | -88.9794
8.9321 | -88.9813
8.9325 | -88.9831
8.9329 | -88.985
8.9332 | -88.9875
8.9336 | -88.9905
8.934 | -88.9928
8.9343 | -88.9949
8.9347 | -88.9966
8.935 | -88.9982
8.9354 | -88.9999
8.9358 | -89.002
8.9361 | -89.0039
8.9365 | -89.0056
8.9368 | -89.008
8.9372 | -89.0096
8.9376 | -89.0113
8.9379 | -89.013
8.9383 | -89.0149
8.9387 | -89.0163
8.939 | -89.0186
8.9394 | -89.0209
8.9397 | -89.0225
8.9401 | -89.0242
8.9405 | -89.0257
8.9408 | -89.0277
8.9412 | -89.0303
8.9416 | -89.0318
8.9419 | -89.0331
8.9423 | -89.0344
8.9426 | -89.0358
8.943 | -89.0374
8.9434 | -89.039
8.9437 | -89.041
8.9441 | -89.0431
8.9444 | -89.0449
8.9448 | -89.0467
8.9452 | -89.0487
8.9455 | -89.0506
8.9459 | -89.052
8.9463 | -89.0536
8.9466 | -89.0552
8.947 | -89.0572
8.9473 | -89.0592
8.9477 | -89.0607
8.9481 | -89.0624
8.9484 | -89.064
8.9488 | -89.0666
8.9491 | -89.0701
8.9495 | -89.0726
8.9499 | -89.0738
8.9502 | -89.0754
8.9506 | -89.0776
8.951 | -89.08
8.9513 | -89.0822
8.9517 | -89.0849
8.952 | -89.0874
8.9524 | -89.0896
8.9528 | -89.092
8.9531 | -89.0943
8.9535 | -89.096
8.9539 | -89.0981
8.9542 | -89.0999
8.9546 | -89.1017
8.9549 | -89.1032
8.9553 | -89.1047
8.9557 | -89.1065
8.956 | -89.1083
8.9564 | -89.1102
8.9567 | -89.1116
8.9571 | -89.1139
8.9575 | -89.1168
8.9578 | -89.1192
8.9582 | -89.1217
8.9586 | -89.1234
8.9589 | -89.1248
8.9593 | -89.1272
8.9596 | -89.1297
8.96 | -89.1316
8.9604 | -89.134
8.9607 | -89.1356
8.9611 | -89.1381
8.9614 | -89.1399
8.9618 | -89.1415
8.9622 | -89.1434
8.9625 | -89.1453
8.9629 | -89.1469
8.9633 | -89.1488
8.9636 | -89.1513
8.964 | -89.1533
8.9643 | -89.1548
8.9647 | -89.1566
8.9651 | -89.1591
8.9654 | -89.1614
8.9658 | -89.1631
8.9662 | -89.1653
8.9665 | -89.1673
8.9669 | -89.169
8.9672 | -89.1704
8.9676 | -89.172
8.968 | -89.1742
8.9683 | -89.1769
8.9687 | -89.179
8.969 | -89.1807
8.9694 | -89.1827
8.9698 | -89.1844
8.9701 | -89.1854
8.9705 | -89.1873
8.9709 | -89.1893
8.9712 | -89.1903
8.9716 | -89.1918
8.9719 | -89.194
8.9723 | -89.1961
8.9727 | -89.198
8.973 | -89.2005
8.9734 | -89.2023
8.9738 | -89.204
8.9741 | -89.2059
8.9745 | -89.2076
8.9748 | -89.209
8.9752 | -89.2108
8.9756 | -89.2129
8.9759 | -89.2146
8.9763 | -89.2163
8.9766 | -89.218
8.977 | -89.2198
8.9774 | -89.2214
8.9777 | -89.2241
8.9781 | -89.2262
8.9785 | -89.2272
8.9788 | -89.2284
8.9792 | -89.229
8.9795 | -89.2309
8.9799 | -89.2327
8.9803 | -89.2343
8.9806 | -89.2357
8.981 | -89.2369
8.9813 | -89.2379
8.9817 | -89.239
8.9821 | -89.2404
8.9824 | -89.242
8.9828 | -89.2437
8.9832 | -89.2455
8.9835 | -89.2472
8.9839 | -89.2483
8.9842 | -89.25
8.9846 | -89.2518
8.985 | -89.2532
8.9853 | -89.2546
8.9857 | -89.2557
8.9861 | -89.2572
8.9864 | -89.2588
8.9868 | -89.26
8.9871 | -89.2614
8.9875 | -89.263
8.9879 | -89.2648
8.9882 | -89.2666
8.9886 | -89.2687
8.9889 | -89.2708
8.9893 | -89.2733
8.9897 | -89.2754
8.99 | -89.2769
8.9904 | -89.2781
8.9908 | -89.2797
8.9911 | -89.281
8.9915 | -89.2823
8.9918 | -89.284
8.9922 | -89.2855
8.9926 | -89.2866
8.9929 | -89.2877
8.9933 | -89.2896
8.9937 | -89.2921
8.994 | -89.2944
8.9944 | -89.2958
8.9947 | -89.2972
8.9951 | -89.2987
8.9955 | -89.3002
8.9958 | -89.3019
8.9962 | -89.3037
8.9965 | -89.3059
8.9969 | -89.3092
8.9973 | -89.3123
8.9976 | -89.3141
8.998 | -89.3156
8.9984 | -89.3172
8.9987 | -89.3196
8.9991 | -89.3221
8.9994 | -89.3236
8.9998 | -89.3244
9.0002 | -89.3249
9.0005 | -89.326
9.0009 | -89.3276
9.0012 | -89.3294
9.0016 | -89.331
9.002 | -89.3325
9.0023 | -89.3346
9.0027 | -89.3356
9.0031 | -89.3365
9.0034 | -89.338
9.0038 | -89.3394
9.0041 | -89.3408
9.0045 | -89.3421
9.0049 | -89.3434
9.0052 | -89.3451
9.0056 | -89.3469
9.006 | -89.3491
9.0063 | -89.3519
9.0067 | -89.3538
9.007 | -89.3555
9.0074 | -89.3573
9.0078 | -89.3591
9.0081 | -89.3605
9.0085 | -89.3613
9.0088 | -89.3627
9.0092 | -89.3643
9.0096 | -89.3661
9.0099 | -89.3671
9.0103 | -89.3677
9.0107 | -89.3692
9.011 | -89.3703
9.0114 | -89.3715
9.0117 | -89.373
9.0121 | -89.3749
9.0125 | -89.3768
9.0128 | -89.3783
9.0132 | -89.3796
9.0135 | -89.3806
9.0139 | -89.382
9.0143 | -89.3833
9.0146 | -89.3847
9.015 | -89.3861
9.0154 | -89.3874
9.0157 | -89.3888
9.0161 | -89.3899
9.0164 | -89.391
9.0168 | -89.3931
9.0172 | -89.3949
9.0175 | -89.3964
9.0179 | -89.3976
9.0183 | -89.3992
9.0186 | -89.4013
9.019 | -89.4026
9.0193 | -89.4041
9.0197 | -89.4055
9.0201 | -89.4069
9.0204 | -89.4083
9.0208 | -89.4089
9.0211 | -89.4085
9.0215 | -89.4094
9.0219 | -89.4109
9.0222 | -89.4122
9.0226 | -89.4128
9.023 | -89.4134
9.0233 | -89.4151
9.0237 | -89.4172
9.024 | -89.4197
9.0244 | -89.4218
9.0248 | -89.4234
9.0251 | -89.425
9.0255 | -89.4266
9.0259 | -89.4279
9.0262 | -89.4296
9.0266 | -89.4314
9.0269 | -89.433
9.0273 | -89.4339
9.0277 | -89.4349
9.028 | -89.4363
9.0284 | -89.4389
9.0287 | -89.4409
9.0291 | -89.4424
9.0295 | -89.4438
9.0298 | -89.4447
9.0302 | -89.446
9.0306 | -89.4473
9.0309 | -89.4483
9.0313 | -89.4494
9.0316 | -89.4507
9.032 | -89.4521
9.0324 | -89.4535
9.0327 | -89.4548
9.0331 | -89.4562
9.0334 | -89.4577
9.0338 | -89.4587
9.0342 | -89.4603
9.0345 | -89.4622
9.0349 | -89.4641
9.0353 | -89.4657
9.0356 | -89.4668
9.036 | -89.468
9.0363 | -89.4696
9.0367 | -89.4717
9.0371 | -89.4729
9.0374 | -89.4742
9.0378 | -89.4755
9.0382 | -89.477
9.0385 | -89.4787
9.0389 | -89.4802
9.0392 | -89.4819
9.0396 | -89.4841
9.04 | -89.4853
9.0403 | -89.4875
9.0407 | -89.4889
9.041 | -89.4901
9.0414 | -89.4916
9.0418 | -89.4936
9.0421 | -89.4954
9.0425 | -89.4968
9.0429 | -89.4981
9.0432 | -89.4995
9.0436 | -89.501
9.0439 | -89.5021
9.0443 | -89.5034
9.0447 | -89.5053
9.045 | -89.5071
9.0454 | -89.5085
9.0457 | -89.5095
9.0461 | -89.5104
9.0465 | -89.5117
9.0468 | -89.5132
9.0472 | -89.5144
9.0476 | -89.5157
9.0479 | -89.5167
9.0483 | -89.5187
9.0486 | -89.5199
9.049 | -89.5206
9.0494 | -89.5219
9.0497 | -89.5236
9.0501 | -89.525
9.0505 | -89.5262
9.0508 | -89.527
9.0512 | -89.5282
9.0515 | -89.5295
9.0519 | -89.5312
9.0523 | -89.5327
9.0526 | -89.5339
9.053 | -89.5352
9.0533 | -89.5365
9.0537 | -89.5379
9.0541 | -89.5399
9.0544 | -89.5414
9.0548 | -89.5423
9.0552 | -89.5445
9.0555 | -89.5465
9.0559 | -89.5478
9.0562 | -89.549
9.0566 | -89.55
9.057 | -89.5504
9.0573 | -89.5507
9.0577 | -89.552
9.0581 | -89.5534
9.0584 | -89.5548
9.0588 | -89.5562
9.0591 | -89.5573
9.0595 | -89.559
9.0599 | -89.5607
9.0602 | -89.5631
9.0606 | -89.5644
9.0609 | -89.567
9.0613 | -89.5683
9.0617 | -89.5695
9.062 | -89.5702
9.0624 | -89.5707
9.0628 | -89.5715
9.0631 | -89.5726
9.0635 | -89.574
9.0638 | -89.5749
9.0642 | -89.5758
9.0646 | -89.5775
9.0649 | -89.5809
9.0653 | -89.5828
9.0656 | -89.584
9.066 | -89.5858
9.0664 | -89.587
9.0667 | -89.5877
9.0671 | -89.5888
9.0675 | -89.5902
9.0678 | -89.5912
9.0682 | -89.592
9.0685 | -89.5928
9.0689 | -89.5941
9.0693 | -89.5955
9.0696 | -89.5966
9.07 | -89.5978
9.0704 | -89.5994
9.0707 | -89.6004
9.0711 | -89.6013
9.0714 | -89.6022
9.0718 | -89.6031
9.0722 | -89.604
9.0725 | -89.6052
9.0729 | -89.6068
9.0732 | -89.6079
9.0736 | -89.6088
9.074 | -89.6096
9.0743 | -89.6101
9.0747 | -89.6111
9.0751 | -89.6126
9.0754 | -89.6136
9.0758 | -89.6147
9.0761 | -89.6157
9.0765 | -89.6165
9.0769 | -89.6172
9.0772 | -89.6183
9.0776 | -89.6197
9.0779 | -89.6212
9.0783 | -89.6222
9.0787 | -89.6231
9.079 | -89.6241
9.0794 | -89.6252
9.0798 | -89.6262
9.0801 | -89.6272
9.0805 | -89.6285
9.0808 | -89.6305
9.0812 | -89.6317
9.0816 | -89.6323
9.0819 | -89.6331
9.0823 | -89.6344
9.0827 | -89.6358
9.083 | -89.6374
9.0834 | -89.6387
9.0837 | -89.6394
9.0841 | -89.6394
9.0845 | -89.6396
9.0848 | -89.6404
9.0852 | -89.6415
9.0855 | -89.6432
9.0859 | -89.6445
9.0863 | -89.6454
9.0866 | -89.6463
9.087 | -89.6474
9.0874 | -89.6483
9.0877 | -89.6491
9.0881 | -89.6497
9.0884 | -89.6498
9.0888 | -89.6506
9.0892 | -89.6515
9.0895 | -89.653
9.0899 | -89.6552
9.0903 | -89.6567
9.0906 | -89.6575
9.091 | -89.6582
9.0913 | -89.6591
9.0917 | -89.66
9.0921 | -89.6611
9.0924 | -89.6621
9.0928 | -89.6638
9.0931 | -89.6654
9.0935 | -89.6665
9.0939 | -89.6677
9.0942 | -89.6689
9.0946 | -89.6698
9.095 | -89.6708
9.0953 | -89.6709
9.0957 | -89.6716
9.096 | -89.6725
9.0964 | -89.6739
9.0968 | -89.6752
9.0971 | -89.6763
9.0975 | -89.677
9.0978 | -89.6784
9.0982 | -89.6798
9.0986 | -89.6807
9.0989 | -89.6813
9.0993 | -89.6835
9.0997 | -89.6851
9.1 | -89.6862
9.1004 | -89.6872
9.1007 | -89.6879
9.1011 | -89.689
9.1015 | -89.6901
9.1018 | -89.6916
9.1022 | -89.6936
9.1026 | -89.6943
9.1029 | -89.6957
9.1033 | -89.6968
9.1036 | -89.6977
9.104 | -89.6984
9.1044 | -89.6991
9.1047 | -89.7
9.1051 | -89.7009
9.1054 | -89.7016
9.1058 | -89.7022
9.1062 | -89.7024
9.1065 | -89.703
9.1069 | -89.7039
9.1073 | -89.7041
9.1076 | -89.7049
9.108 | -89.7057
9.1083 | -89.7069
9.1087 | -89.7079
9.1091 | -89.7085
9.1094 | -89.7095
9.1098 | -89.7102
9.1102 | -89.7112
9.1105 | -89.7123
9.1109 | -89.7132
9.1112 | -89.714
9.1116 | -89.7148
9.112 | -89.7146
9.1123 | -89.7137
9.1127 | -89.7142
9.113 | -89.7146
9.1134 | -89.7151
9.1138 | -89.7166
9.1141 | -89.7179
9.1145 | -89.7188
9.1149 | -89.7195
9.1152 | -89.7204
9.1156 | -89.7205
9.1159 | -89.7208
9.1163 | -89.7219
9.1167 | -89.7226
9.117 | -89.7231
9.1174 | -89.7238
9.1177 | -89.7249
9.1181 | -89.725
9.1185 | -89.7255
9.1188 | -89.7269
9.1192 | -89.7283
9.1196 | -89.7284
9.1199 | -89.7293
9.1203 | -89.73
9.1206 | -89.7304
9.121 | -89.7313
9.1214 | -89.7323
9.1217 | -89.7335
9.1221 | -89.7344
9.1225 | -89.7357
9.1228 | -89.7377
9.1232 | -89.7393
9.1235 | -89.7403
9.1239 | -89.7414
9.1243 | -89.7428
9.1246 | -89.7438
9.125 | -89.7446
9.1253 | -89.7454
9.1257 | -89.7464
9.1261 | -89.7471
9.1264 | -89.7479
9.1268 | -89.7488
9.1272 | -89.7497
9.1275 | -89.7508
9.1279 | -89.7514
9.1282 | -89.7522
9.1286 | -89.7531
9.129 | -89.7538
9.1293 | -89.7545
9.1297 | -89.7552
9.13 | -89.7557
9.1304 | -89.7574
9.1308 | -89.7588
9.1311 | -89.7598
9.1315 | -89.761
9.1319 | -89.7617
9.1322 | -89.7619
9.1326 | -89.7628
9.1329 | -89.7637
9.1333 | -89.7645
9.1337 | -89.7651
9.134 | -89.7652
9.1344 | -89.7658
9.1348 | -89.7663
9.1351 | -89.7669
9.1355 | -89.768
9.1358 | -89.768
9.1362 | -89.7675
9.1366 | -89.7673
9.1369 | -89.7683
9.1373 | -89.7698
9.1376 | -89.7705
9.138 | -89.7712
9.1384 | -89.7723
9.1387 | -89.7732
9.1391 | -89.774
9.1395 | -89.7744
9.1398 | -89.7745
9.1402 | -89.7752
9.1405 | -89.7756
9.1409 | -89.7762
9.1413 | -89.7768
9.1416 | -89.7774
9.142 | -89.778
9.1424 | -89.7788
9.1427 | -89.7789
9.1431 | -89.779
9.1434 | -89.7792
9.1438 | -89.7801
9.1442 | -89.7808
9.1445 | -89.7817
9.1449 | -89.7823
9.1452 | -89.7832
9.1456 | -89.7846
9.146 | -89.7858
9.1463 | -89.7869
9.1467 | -89.7872
9.1471 | -89.7872
9.1474 | -89.7874
9.1478 | -89.7881
9.1481 | -89.7887
9.1485 | -89.7893
9.1489 | -89.7908
9.1492 | -89.7921
9.1496 | -89.7922
9.1499 | -89.7936
9.1503 | -89.7947
9.1507 | -89.7957
9.151 | -89.7966
9.1514 | -89.7978
9.1518 | -89.7996
9.1521 | -89.8005
9.1525 | -89.8012
9.1528 | -89.8033
9.1532 | -89.8061
9.1536 | -89.808
9.1539 | -89.8098
9.1543 | -89.8101
9.1547 | -89.8094
9.155 | -89.8101
9.1554 | -89.8108
9.1557 | -89.812
9.1561 | -89.8137
9.1565 | -89.8153
9.1568 | -89.8162
9.1572 | -89.8171
9.1575 | -89.8188
9.1579 | -89.8202
9.1583 | -89.821
9.1586 | -89.8209
9.159 | -89.8213
9.1594 | -89.8223
9.1597 | -89.8239
9.1601 | -89.8248
9.1604 | -89.8253
9.1608 | -89.8258
9.1612 | -89.826
9.1615 | -89.8263
9.1619 | -89.8272
9.1623 | -89.828
9.1626 | -89.8284
9.163 | -89.8284
9.1633 | -89.8291
9.1637 | -89.8291
9.1641 | -89.8294
9.1644 | -89.8299
9.1648 | -89.8304
9.1651 | -89.8302
9.1655 | -89.8311
9.1659 | -89.8329
9.1662 | -89.8342
9.1666 | -89.8345
9.167 | -89.8354
9.1673 | -89.8365
9.1677 | -89.8367
9.168 | -89.8369
9.1684 | -89.8377
9.1688 | -89.8394
9.1691 | -89.8401
9.1695 | -89.8405
9.1698 | -89.8405
9.1702 | -89.841
9.1706 | -89.8412
9.1709 | -89.8416
9.1713 | -89.8426
9.1717 | -89.8434
9.172 | -89.8439
9.1724 | -89.8451
9.1727 | -89.8459
9.1731 | -89.8462
9.1735 | -89.8465
9.1738 | -89.8473
9.1742 | -89.848
9.1746 | -89.8495
9.1749 | -89.8519
9.1753 | -89.8545
9.1756 | -89.8553
9.176 | -89.8572
9.1764 | -89.8576
9.1767 | -89.8583
9.1771 | -89.859
9.1774 | -89.8593
9.1778 | -89.8597
9.1782 | -89.8611
9.1785 | -89.8628
9.1789 | -89.8642
9.1793 | -89.8648
9.1796 | -89.8654
9.18 | -89.8659
9.1803 | -89.8661
9.1807 | -89.8667
9.1811 | -89.8673
9.1814 | -89.8682
9.1818 | -89.8687
9.1821 | -89.8689
9.1825 | -89.8694
9.1829 | -89.8704
9.1832 | -89.8712
9.1836 | -89.8716
9.184 | -89.8723
9.1843 | -89.8727
9.1847 | -89.8732
9.185 | -89.8738
9.1854 | -89.8748
9.1858 | -89.8756
9.1861 | -89.8765
9.1865 | -89.877
9.1869 | -89.8773
9.1872 | -89.8772
9.1876 | -89.8779
9.1879 | -89.8794
9.1883 | -89.8798
9.1887 | -89.8804
9.189 | -89.8808
9.1894 | -89.8808
9.1897 | -89.8808
9.1901 | -89.881
9.1905 | -89.8817
9.1908 | -89.8823
9.1912 | -89.8823
9.1916 | -89.8826
9.1919 | -89.8829
9.1923 | -89.8831
9.1926 | -89.8837
9.193 | -89.8845
9.1934 | -89.8858
9.1937 | -89.8872
9.1941 | -89.8884
9.1944 | -89.8893
9.1948 | -89.8902
9.1952 | -89.8909
9.1955 | -89.8915
9.1959 | -89.8921
9.1963 | -89.8929
9.1966 | -89.8937
9.197 | -89.8941
9.1973 | -89.8946
9.1977 | -89.8954
9.1981 | -89.8959
9.1984 | -89.8962
9.1988 | -89.8966
9.1992 | -89.8975
9.1995 | -89.8984
9.1999 | -89.8986
9.2002 | -89.8989
9.2006 | -89.8994
9.201 | -89.9
9.2013 | -89.9004
9.2017 | -89.9015
9.202 | -89.9032
9.2024 | -89.9042
9.2028 | -89.9044
9.2031 | -89.9047
9.2035 | -89.9047
9.2039 | -89.9043
9.2042 | -89.9043
9.2046 | -89.9046
9.2049 | -89.9051
9.2053 | -89.906
9.2057 | -89.9072
9.206 | -89.9078
9.2064 | -89.9079
9.2068 | -89.9086
9.2071 | -89.9089
9.2075 | -89.9098
9.2078 | -89.9104
9.2082 | -89.91
9.2086 | -89.9106
9.2089 | -89.9114
9.2093 | -89.9119
9.2096 | -89.9124
9.21 | -89.9137
9.2104 | -89.916
9.2107 | -89.9184
9.2111 | -89.9192
9.2115 | -89.9194
9.2118 | -89.9194
9.2122 | -89.9197
9.2125 | -89.9202
9.2129 | -89.9207
9.2133 | -89.9214
9.2136 | -89.9217
9.214 | -89.9218
9.2143 | -89.9221
9.2147 | -89.9224
9.2151 | -89.9225
9.2154 | -89.9227
9.2158 | -89.9234
9.2162 | -89.9243
9.2165 | -89.9249
9.2169 | -89.9253
9.2172 | -89.926
9.2176 | -89.9264
9.218 | -89.9267
9.2183 | -89.9266
9.2187 | -89.9267
9.2191 | -89.9271
9.2194 | -89.9274
9.2198 | -89.9287
9.2201 | -89.9298
9.2205 | -89.9305
9.2209 | -89.9315
9.2212 | -89.9319
9.2216 | -89.9325
9.2219 | -89.933
9.2223 | -89.9338
9.2227 | -89.9344
9.223 | -89.9346
9.2234 | -89.9347
9.2238 | -89.9347
9.2241 | -89.9352
9.2245 | -89.9363
9.2248 | -89.937
9.2252 | -89.9374
9.2256 | -89.9383
9.2259 | -89.9388
9.2263 | -89.9391
9.2266 | -89.9395
9.227 | -89.9397
9.2274 | -89.9402
9.2277 | -89.9411
9.2281 | -89.9414
9.2285 | -89.9413
9.2288 | -89.9417
9.2292 | -89.9419
9.2295 | -89.9427
9.2299 | -89.9431
9.2303 | -89.9437
9.2306 | -89.9443
9.231 | -89.945
9.2314 | -89.9455
9.2317 | -89.946
9.2321 | -89.9464
9.2324 | -89.9468
9.2328 | -89.9476
9.2332 | -89.9488
9.2335 | -89.9489
9.2339 | -89.949
9.2342 | -89.9494
9.2346 | -89.9501
9.235 | -89.9508
9.2353 | -89.9512
9.2357 | -89.9516
9.2361 | -89.9523
9.2364 | -89.953
9.2368 | -89.9534
9.2371 | -89.9541
9.2375 | -89.9548
9.2379 | -89.9553
9.2382 | -89.9563
9.2386 | -89.9568
9.239 | -89.9572
9.2393 | -89.9571
9.2397 | -89.9576
9.24 | -89.9587
9.2404 | -89.9593
9.2408 | -89.9598
9.2411 | -89.9605
9.2415 | -89.9613
9.2418 | -89.9625
9.2422 | -89.9629
9.2426 | -89.9636
9.2429 | -89.9647
9.2433 | -89.9656
9.2437 | -89.9662
9.244 | -89.9667
9.2444 | -89.9669
9.2447 | -89.9664
9.2451 | -89.9664
9.2455 | -89.9669
9.2458 | -89.9675
9.2462 | -89.9679
9.2465 | -89.9682
9.2469 | -89.9692
9.2473 | -89.9703
9.2476 | -89.9707
9.248 | -89.9712
9.2484 | -89.9719
9.2487 | -89.9714
9.2491 | -89.9729
9.2494 | -89.9737
9.2498 | -89.9743
9.2502 | -89.9748
9.2505 | -89.9753
9.2509 | -89.976
9.2513 | -89.9765
9.2516 | -89.9771
9.252 | -89.9783
9.2523 | -89.9793
9.2527 | -89.9796
9.2531 | -89.9799
9.2534 | -89.9804
9.2538 | -89.9805
9.2541 | -89.9802
9.2545 | -89.9806
9.2549 | -89.9811
9.2552 | -89.9815
9.2556 | -89.9822
9.256 | -89.9832
9.2563 | -89.9837
9.2567 | -89.9842
9.257 | -89.9846
9.2574 | -89.9855
9.2578 | -89.9867
9.2581 | -89.9877
9.2585 | -89.9883
9.2589 | -89.9887
9.2592 | -89.9896
9.2596 | -89.9899
9.2599 | -89.9904
9.2603 | -89.991
9.2607 | -89.9916
9.261 | -89.9921
9.2614 | -89.9927
9.2617 | -89.9932
9.2621 | -89.9941
9.2625 | -89.995
9.2628 | -89.9958
9.2632 | -89.9964
9.2636 | -89.997
9.2639 | -89.9973
9.2643 | -89.9978
9.2646 | -89.9987
9.265 | -89.9992
9.2654 | -89.9993
9.2657 | -89.9999
9.2661 | -90.0007
9.2664 | -90.0019
9.2668 | -90.0027
9.2672 | -90.0035
9.2675 | -90.0046
9.2679 | -90.0051
9.2683 | -90.0058
9.2686 | -90.0067
9.269 | -90.0072
9.2693 | -90.0079
9.2697 | -90.0081
9.2701 | -90.0088
9.2704 | -90.0095
9.2708 | -90.0101
9.2712 | -90.0108
9.2715 | -90.0114
9.2719 | -90.012
9.2722 | -90.0124
9.2726 | -90.013
9.273 | -90.0136
9.2733 | -90.0136
9.2737 | -90.0137
9.274 | -90.0144
9.2744 | -90.0156
9.2748 | -90.0172
9.2751 | -90.0179
9.2755 | -90.0182
9.2759 | -90.0188
9.2762 | -90.0199
9.2766 | -90.0203
9.2769 | -90.0206
9.2773 | -90.0206
9.2777 | -90.0211
9.278 | -90.0225
9.2784 | -90.0234
9.2788 | -90.024
9.2791 | -90.0245
9.2795 | -90.0255
9.2798 | -90.0261
9.2802 | -90.0265
9.2806 | -90.0268
9.2809 | -90.0271
9.2813 | -90.0273
9.2816 | -90.0278
9.282 | -90.0285
9.2824 | -90.0295
9.2827 | -90.03
9.2831 | -90.0305
9.2835 | -90.0311
9.2838 | -90.0316
9.2842 | -90.0322
9.2845 | -90.0328
9.2849 | -90.0333
9.2853 | -90.0339
9.2856 | -90.0344
9.286 | -90.0357
9.2863 | -90.037
9.2867 | -90.0377
9.2871 | -90.0393
9.2874 | -90.0399
9.2878 | -90.0405
9.2882 | -90.0412
9.2885 | -90.0418
9.2889 | -90.0427
9.2892 | -90.0439
9.2896 | -90.0446
9.29 | -90.0451
9.2903 | -90.0464
9.2907 | -90.0469
9.2911 | -90.0473
9.2914 | -90.0479
9.2918 | -90.048
9.2921 | -90.0485
9.2925 | -90.0492
9.2929 | -90.0496
9.2932 | -90.0504
9.2936 | -90.0519
9.2939 | -90.0533
9.2943 | -90.0537
9.2947 | -90.0544
9.295 | -90.055
9.2954 | -90.0557
9.2958 | -90.0563
9.2961 | -90.0572
9.2965 | -90.058
9.2968 | -90.0585
9.2972 | -90.0592
9.2976 | -90.06
9.2979 | -90.061
9.2983 | -90.0616
9.2986 | -90.0627
9.299 | -90.063
9.2994 | -90.0641
9.2997 | -90.0647
9.3001 | -90.0653
9.3005 | -90.0655
9.3008 | -90.0664
9.3012 | -90.067
9.3015 | -90.0676
9.3019 | -90.0687
9.3023 | -90.0693
9.3026 | -90.0699
9.303 | -90.0704
9.3034 | -90.0711
9.3037 | -90.0719
9.3041 | -90.0726
9.3044 | -90.0732
9.3048 | -90.0738
9.3052 | -90.0741
9.3055 | -90.0746
9.3059 | -90.0752
9.3062 | -90.0757
9.3066 | -90.0764
9.307 | -90.0768
9.3073 | -90.0771
9.3077 | -90.0777
9.3081 | -90.0783
9.3084 | -90.0792
9.3088 | -90.08
9.3091 | -90.0805
9.3095 | -90.0816
9.3099 | -90.0828
9.3102 | -90.0834
9.3106 | -90.084
9.311 | -90.0845
9.3113 | -90.0853
9.3117 | -90.086
9.312 | -90.0871
9.3124 | -90.088
9.3128 | -90.0888
9.3131 | -90.0893
9.3135 | -90.0898
9.3138 | -90.0889
9.3142 | -90.0898
9.3146 | -90.0917
9.3149 | -90.0932
9.3153 | -90.0945
9.3157 | -90.0957
9.316 | -90.0964
9.3164 | -90.0979
9.3167 | -90.0992
9.3171 | -90.1
9.3175 | -90.1008
9.3178 | -90.1015
9.3182 | -90.1022
9.3185 | -90.1032
9.3189 | -90.1037
9.3193 | -90.1046
9.3196 | -90.1058
9.32 | -90.1069
9.3204 | -90.1077
9.3207 | -90.1087
9.3211 | -90.1095
9.3214 | -90.1097
9.3218 | -90.1096
9.3222 | -90.1105
9.3225 | -90.1114
9.3229 | -90.112
9.3233 | -90.1129
9.3236 | -90.1141
9.324 | -90.115
9.3243 | -90.1157
9.3247 | -90.1164
9.3251 | -90.117
9.3254 | -90.1178
9.3258 | -90.1192
9.3261 | -90.1198
9.3265 | -90.1201
9.3269 | -90.1208
9.3272 | -90.1212
9.3276 | -90.1218
9.328 | -90.1224
9.3283 | -90.123
9.3287 | -90.1235
9.329 | -90.1241
9.3294 | -90.1248
9.3298 | -90.1258
9.3301 | -90.1267
9.3305 | -90.1275
9.3308 | -90.1282
9.3312 | -90.1291
9.3316 | -90.1298
9.3319 | -90.1309
9.3323 | -90.1322
9.3327 | -90.1334
9.333 | -90.1339
9.3334 | -90.1346
9.3337 | -90.1359
9.3341 | -90.1374
9.3345 | -90.1382
9.3348 | -90.1389
9.3352 | -90.1396
9.3356 | -90.1403
9.3359 | -90.1408
9.3363 | -90.1415
9.3366 | -90.1426
9.337 | -90.1449
9.3374 | -90.1458
9.3377 | -90.1467
9.3381 | -90.1472
9.3384 | -90.148
9.3388 | -90.1486
9.3392 | -90.1491
9.3395 | -90.1497
9.3399 | -90.1501
9.3403 | -90.1503
9.3406 | -90.1505
9.341 | -90.1512
9.3413 | -90.1522
9.3417 | -90.1532
9.3421 | -90.1538
9.3424 | -90.1549
9.3428 | -90.1556
9.3431 | -90.1561
9.3435 | -90.157
9.3439 | -90.1577
9.3442 | -90.1587
9.3446 | -90.1599
9.345 | -90.1618
9.3453 | -90.1637
9.3457 | -90.1656
9.346 | -90.1675
9.3464 | -90.1705
9.3468 | -90.1723
9.3471 | -90.1736
9.3475 | -90.1744
9.3479 | -90.1748
9.3482 | -90.1752
9.3486 | -90.176
9.3489 | -90.1769
9.3493 | -90.1776
9.3497 | -90.1784
9.35 | -90.1792
9.3504 | -90.18
9.3507 | -90.1809
9.3511 | -90.1818
9.3515 | -90.1825
9.3518 | -90.1839
9.3522 | -90.1864
9.3526 | -90.1875
9.3529 | -90.1886
9.3533 | -90.1894
9.3536 | -90.1902
9.354 | -90.1911
9.3544 | -90.1918
9.3547 | -90.1926
9.3551 | -90.1935
9.3555 | -90.1941
9.3558 | -90.1946
9.3562 | -90.1959
9.3565 | -90.1972
9.3569 | -90.1984
9.3573 | -90.1993
9.3576 | -90.2003
9.358 | -90.2011
9.3583 | -90.202
9.3587 | -90.2021
9.3591 | -90.2026
9.3594 | -90.2044
9.3598 | -90.2056
9.3602 | -90.2068
9.3605 | -90.2079
9.3609 | -90.2084
9.3612 | -90.2088
9.3616 | -90.2097
9.362 | -90.2106
9.3623 | -90.2115
9.3627 | -90.2123
9.363 | -90.2132
9.3634 | -90.2146
9.3638 | -90.216
9.3641 | -90.2172
9.3645 | -90.2179
9.3649 | -90.2184
9.3652 | -90.2186
9.3656 | -90.2195
9.3659 | -90.2203
9.3663 | -90.221
9.3667 | -90.2221
9.367 | -90.2234
9.3674 | -90.2246
9.3678 | -90.2245
9.3681 | -90.2247
9.3685 | -90.225
9.3688 | -90.2257
9.3692 | -90.2263
9.3696 | -90.2277
9.3699 | -90.229
9.3703 | -90.2305
9.3706 | -90.2337
9.371 | -90.2356
9.3714 | -90.2363
9.3717 | -90.2372
9.3721 | -90.238
9.3725 | -90.2386
9.3728 | -90.2392
9.3732 | -90.2399
9.3735 | -90.2407
9.3739 | -90.2419
9.3743 | -90.2422
9.3746 | -90.2435
9.375 | -90.2444
9.3753 | -90.2451
9.3757 | -90.2461
9.3761 | -90.2483
9.3764 | -90.2505
9.3768 | -90.2514
9.3772 | -90.252
9.3775 | -90.2528
9.3779 | -90.2543
9.3782 | -90.2557
9.3786 | -90.2571
9.379 | -90.2583
9.3793 | -90.2592
9.3797 | -90.2603
9.3801 | -90.2618
9.3804 | -90.2631
9.3808 | -90.2632
9.3811 | -90.2637
9.3815 | -90.2649
9.3819 | -90.2659
9.3822 | -90.2667
9.3826 | -90.2676
9.3829 | -90.2686
9.3833 | -90.2693
9.3837 | -90.2696
9.384 | -90.2715
9.3844 | -90.273
9.3848 | -90.2742
9.3851 | -90.2755
9.3855 | -90.2765
9.3858 | -90.2775
9.3862 | -90.2785
9.3866 | -90.2793
9.3869 | -90.2802
9.3873 | -90.2812
9.3877 | -90.2827
9.388 | -90.2839
9.3884 | -90.2848
9.3887 | -90.286
9.3891 | -90.2883
9.3895 | -90.2903
9.3898 | -90.2918
9.3902 | -90.2934
9.3905 | -90.2953
9.3909 | -90.2971
9.3913 | -90.2979
9.3916 | -90.2992
9.392 | -90.3007
9.3924 | -90.3016
9.3927 | -90.3025
9.3931 | -90.3045
9.3934 | -90.3065
9.3938 | -90.3075
9.3942 | -90.3081
9.3945 | -90.3088
9.3949 | -90.3098
9.3953 | -90.3105
9.3956 | -90.3116
9.396 | -90.3125
9.3963 | -90.3132
9.3967 | -90.3137
9.3971 | -90.3144
9.3974 | -90.3155
9.3978 | -90.3163
9.3981 | -90.3172
9.3985 | -90.3183
9.3989 | -90.3194
9.3992 | -90.3204
9.3996 | -90.3214
9.4 | -90.3224
9.4003 | -90.3234
9.4007 | -90.3244
9.401 | -90.3253
9.4014 | -90.3264
9.4018 | -90.3274
9.4021 | -90.3288
9.4025 | -90.3302
9.4028 | -90.331
9.4032 | -90.3311
9.4036 | -90.3323
9.4039 | -90.334
9.4043 | -90.3348
9.4047 | -90.3356
9.405 | -90.3366
9.4054 | -90.3376
9.4057 | -90.3383
9.4061 | -90.3392
9.4065 | -90.3402
9.4068 | -90.3412
9.4072 | -90.3422
9.4076 | -90.3432
9.4079 | -90.3442
9.4083 | -90.3455
9.4086 | -90.3468
9.409 | -90.3479
9.4094 | -90.3492
9.4097 | -90.3515
9.4101 | -90.3529
9.4104 | -90.3541
9.4108 | -90.3555
9.4112 | -90.3564
9.4115 | -90.3569
9.4119 | -90.3581
9.4123 | -90.3594
9.4126 | -90.3602
9.413 | -90.3608
9.4133 | -90.3617
9.4137 | -90.3627
9.4141 | -90.3638
9.4144 | -90.3649
9.4148 | -90.3661
9.4151 | -90.367
9.4155 | -90.3685
9.4159 | -90.37
9.4162 | -90.3718
9.4166 | -90.3727
9.417 | -90.3737
9.4173 | -90.3743
9.4177 | -90.3757
9.418 | -90.3774
9.4184 | -90.3783
9.4188 | -90.3796
9.4191 | -90.3804
9.4195 | -90.3814
9.4199 | -90.3824
9.4202 | -90.3835
9.4206 | -90.3846
9.4209 | -90.3861
9.4213 | -90.3866
9.4217 | -90.3868
9.422 | -90.3876
9.4224 | -90.3885
9.4227 | -90.3895
9.4231 | -90.3909
9.4235 | -90.393
9.4238 | -90.3943
9.4242 | -90.3953
9.4246 | -90.3964
9.4249 | -90.3973
9.4253 | -90.3983
9.4256 | -90.3996
9.426 | -90.4007
9.4264 | -90.402
9.4267 | -90.4032
9.4271 | -90.4038
9.4275 | -90.4035
9.4278 | -90.4043
9.4282 | -90.405
9.4285 | -90.4059
9.4289 | -90.4068
9.4293 | -90.408
9.4296 | -90.4091
9.43 | -90.4098
9.4303 | -90.411
9.4307 | -90.4122
9.4311 | -90.4133
9.4314 | -90.4144
9.4318 | -90.4155
9.4322 | -90.4167
9.4325 | -90.4176
9.4329 | -90.418
9.4332 | -90.4182
9.4336 | -90.4195
9.434 | -90.4207
9.4343 | -90.4217
9.4347 | -90.423
9.435 | -90.4242
9.4354 | -90.4253
9.4358 | -90.4263
9.4361 | -90.4272
9.4365 | -90.4284
9.4369 | -90.4295
9.4372 | -90.4313
9.4376 | -90.434
9.4379 | -90.4359
9.4383 | -90.4368
9.4387 | -90.4382
9.439 | -90.4395
9.4394 | -90.4405
9.4398 | -90.4413
9.4401 | -90.4423
9.4405 | -90.4431
9.4408 | -90.4444
9.4412 | -90.4459
9.4416 | -90.4473
9.4419 | -90.4482
9.4423 | -90.4491
9.4426 | -90.4505
9.443 | -90.4516
9.4434 | -90.4528
9.4437 | -90.4538
9.4441 | -90.4548
9.4445 | -90.4559
9.4448 | -90.4571
9.4452 | -90.4583
9.4455 | -90.4597
9.4459 | -90.4609
9.4463 | -90.462
9.4466 | -90.4632
9.447 | -90.4647
9.4473 | -90.4662
9.4477 | -90.4676
9.4481 | -90.4689
9.4484 | -90.4703
9.4488 | -90.472
9.4492 | -90.4732
9.4495 | -90.4743
9.4499 | -90.4755
9.4502 | -90.4767
9.4506 | -90.4777
9.451 | -90.4783
9.4513 | -90.4799
9.4517 | -90.481
9.4521 | -90.4822
9.4524 | -90.4847
9.4528 | -90.4867
9.4531 | -90.4877
9.4535 | -90.4899
9.4539 | -90.4922
9.4542 | -90.4935
9.4546 | -90.4946
9.4549 | -90.4955
9.4553 | -90.4969
9.4557 | -90.4988
9.456 | -90.4999
9.4564 | -90.501
9.4568 | -90.5023
9.4571 | -90.5037
9.4575 | -90.5052
9.4578 | -90.5057
9.4582 | -90.5064
9.4586 | -90.5078
9.4589 | -90.509
9.4593 | -90.5098
9.4596 | -90.5098
9.46 | -90.5098
9.4604 | -90.5101
9.4607 | -90.5108
9.4611 | -90.5121
9.4615 | -90.5128
9.4618 | -90.5131
9.4622 | -90.5141
9.4625 | -90.5157
9.4629 | -90.5184
9.4633 | -90.5208
9.4636 | -90.5222
9.464 | -90.5233
9.4644 | -90.5248
9.4647 | -90.5259
9.4651 | -90.5271
9.4654 | -90.5282
9.4658 | -90.5292
9.4662 | -90.5303
9.4665 | -90.5317
9.4669 | -90.5329
9.4672 | -90.534
9.4676 | -90.5355
9.468 | -90.5372
9.4683 | -90.5386
9.4687 | -90.5397
9.4691 | -90.5406
9.4694 | -90.542
9.4698 | -90.5434
9.4701 | -90.5452
9.4705 | -90.5466
9.4709 | -90.5476
9.4712 | -90.5483
9.4716 | -90.5488
9.472 | -90.5492
9.4723 | -90.5504
9.4727 | -90.5517
9.473 | -90.5528
9.4734 | -90.554
9.4738 | -90.5555
9.4741 | -90.5568
9.4745 | -90.5579
9.4748 | -90.5591
9.4752 | -90.5602
9.4756 | -90.5613
9.4759 | -90.5625
9.4763 | -90.5637
9.4767 | -90.565
9.477 | -90.5666
9.4774 | -90.5693
9.4777 | -90.5706
9.4781 | -90.5721
9.4785 | -90.5734
9.4788 | -90.5743
9.4792 | -90.5755
9.4795 | -90.5767
9.4799 | -90.5779
9.4803 | -90.5791
9.4806 | -90.5806
9.481 | -90.5819
9.4814 | -90.5834
9.4817 | -90.5845
9.4821 | -90.5842
9.4824 | -90.5846
9.4828 | -90.586
9.4832 | -90.5874
9.4835 | -90.5889
9.4839 | -90.5905
9.4843 | -90.592
9.4846 | -90.5931
9.485 | -90.5943
9.4853 | -90.5957
9.4857 | -90.5969
9.4861 | -90.5984
9.4864 | -90.5998
9.4868 | -90.6014
9.4871 | -90.6033
9.4875 | -90.6043
9.4879 | -90.6057
9.4882 | -90.6074
9.4886 | -90.609
9.489 | -90.6105
9.4893 | -90.6122
9.4897 | -90.6146
9.49 | -90.6162
9.4904 | -90.6174
9.4908 | -90.6187
9.4911 | -90.6199
9.4915 | -90.6216
9.4918 | -90.6231
9.4922 | -90.6246
9.4926 | -90.6257
9.4929 | -90.6269
9.4933 | -90.6284
9.4937 | -90.6296
9.494 | -90.6309
9.4944 | -90.6325
9.4947 | -90.6344
9.4951 | -90.6365
9.4955 | -90.6397
9.4958 | -90.6421
9.4962 | -90.6431
9.4966 | -90.6437
9.4969 | -90.6448
9.4973 | -90.6465
9.4976 | -90.6478
9.498 | -90.649
9.4984 | -90.6505
9.4987 | -90.6525
9.4991 | -90.6538
9.4994 | -90.6549
9.4998 | -90.6559
9.5002 | -90.6572
9.5005 | -90.6586
9.5009 | -90.6604
9.5013 | -90.6621
9.5016 | -90.6632
9.502 | -90.6641
9.5023 | -90.6648
9.5027 | -90.666
9.5031 | -90.6671
9.5034 | -90.6683
9.5038 | -90.6701
9.5042 | -90.6709
9.5045 | -90.6728
9.5049 | -90.6744
9.5052 | -90.6747
9.5056 | -90.6755
9.506 | -90.6768
9.5063 | -90.678
9.5067 | -90.6792
9.507 | -90.6804
9.5074 | -90.6816
9.5078 | -90.6829
9.5081 | -90.6846
9.5085 | -90.6861
9.5089 | -90.6878
9.5092 | -90.6892
9.5096 | -90.6906
9.5099 | -90.6919
9.5103 | -90.6931
9.5107 | -90.6943
9.511 | -90.6951
9.5114 | -90.6965
9.5117 | -90.6979
9.5121 | -90.699
9.5125 | -90.7
9.5128 | -90.7012
9.5132 | -90.7025
9.5136 | -90.7036
9.5139 | -90.7056
9.5143 | -90.7082
9.5146 | -90.7096
9.515 | -90.7109
9.5154 | -90.7129
9.5157 | -90.714
9.5161 | -90.7158
9.5165 | -90.7173
9.5168 | -90.7184
9.5172 | -90.7195
9.5175 | -90.7208
9.5179 | -90.7223
9.5183 | -90.724
9.5186 | -90.7257
9.519 | -90.7271
9.5193 | -90.7283
9.5197 | -90.7294
9.5201 | -90.7306
9.5204 | -90.7314
9.5208 | -90.7325
9.5212 | -90.7337
9.5215 | -90.7349
9.5219 | -90.7361
9.5222 | -90.737
9.5226 | -90.7387
9.523 | -90.7402
9.5233 | -90.7413
9.5237 | -90.7429
9.5241 | -90.7446
9.5244 | -90.7459
9.5248 | -90.7473
9.5251 | -90.7484
9.5255 | -90.7496
9.5259 | -90.7505
9.5262 | -90.7516
9.5266 | -90.7531
9.5269 | -90.7545
9.5273 | -90.7561
9.5277 | -90.7572
9.528 | -90.7584
9.5284 | -90.7596
9.5288 | -90.7604
9.5291 | -90.7614
9.5295 | -90.763
9.5298 | -90.7645
9.5302 | -90.7658
9.5306 | -90.7673
9.5309 | -90.7683
9.5313 | -90.7692
9.5316 | -90.7704
9.532 | -90.7712
9.5324 | -90.772
9.5327 | -90.7735
9.5331 | -90.775
9.5335 | -90.7762
9.5338 | -90.7773
9.5342 | -90.7784
9.5345 | -90.7797
9.5349 | -90.781
9.5353 | -90.7824
9.5356 | -90.7835
9.536 | -90.7847
9.5364 | -90.786
9.5367 | -90.7876
9.5371 | -90.7885
9.5374 | -90.7908
9.5378 | -90.7921
9.5382 | -90.7934
9.5385 | -90.795
9.5389 | -90.7964
9.5392 | -90.7977
9.5396 | -90.7989
9.54 | -90.8006
9.5403 | -90.8019
9.5407 | -90.803
9.5411 | -90.8043
9.5414 | -90.8053
9.5418 | -90.8063
9.5421 | -90.808
9.5425 | -90.8095
9.5429 | -90.8107
9.5432 | -90.8119
9.5436 | -90.8132
9.544 | -90.8144
9.5443 | -90.8152
9.5447 | -90.8163
9.545 | -90.8173
9.5454 | -90.8184
9.5458 | -90.8199
9.5461 | -90.8213
9.5465 | -90.8223
9.5468 | -90.8241
9.5472 | -90.8251
9.5476 | -90.826
9.5479 | -90.8272
9.5483 | -90.8286
9.5487 | -90.8296
9.549 | -90.8303
9.5494 | -90.8312
9.5497 | -90.8326
9.5501 | -90.8342
9.5505 | -90.8358
9.5508 | -90.8373
9.5512 | -90.8385
9.5515 | -90.8394
9.5519 | -90.8401
9.5523 | -90.8414
9.5526 | -90.8426
9.553 | -90.8438
9.5534 | -90.8457
9.5537 | -90.8467
9.5541 | -90.8475
9.5544 | -90.8486
9.5548 | -90.85
9.5552 | -90.8513
9.5555 | -90.8526
9.5559 | -90.8538
9.5563 | -90.8548
9.5566 | -90.8562
9.557 | -90.858
9.5573 | -90.8599
9.5577 | -90.8606
9.5581 | -90.8611
9.5584 | -90.8622
9.5588 | -90.8637
9.5591 | -90.8663
9.5595 | -90.8684
9.5599 | -90.8694
9.5602 | -90.8704
9.5606 | -90.8714
9.561 | -90.8721
9.5613 | -90.8724
9.5617 | -90.8732
9.562 | -90.874
9.5624 | -90.8748
9.5628 | -90.8759
9.5631 | -90.8771
9.5635 | -90.8777
9.5638 | -90.878
9.5642 | -90.879
9.5646 | -90.8801
9.5649 | -90.8814
9.5653 | -90.8829
9.5657 | -90.8839
9.566 | -90.885
9.5664 | -90.8865
9.5667 | -90.8875
9.5671 | -90.8889
9.5675 | -90.8903
9.5678 | -90.8914
9.5682 | -90.8925
9.5686 | -90.8935
9.5689 | -90.8945
9.5693 | -90.8954
9.5696 | -90.8964
9.57 | -90.8975
9.5704 | -90.8987
9.5707 | -90.8998
9.5711 | -90.9011
9.5714 | -90.9034
9.5718 | -90.9068
9.5722 | -90.9075
9.5725 | -90.9102
9.5729 | -90.9113
9.5733 | -90.9128
9.5736 | -90.9145
9.574 | -90.9152
9.5743 | -90.916
9.5747 | -90.917
9.5751 | -90.9181
9.5754 | -90.9192
9.5758 | -90.9201
9.5762 | -90.9208
9.5765 | -90.9215
9.5769 | -90.9223
9.5772 | -90.9233
9.5776 | -90.9248
9.578 | -90.9253
9.5783 | -90.9256
9.5787 | -90.9266
9.579 | -90.928
9.5794 | -90.9291
9.5798 | -90.9303
9.5801 | -90.9315
9.5805 | -90.9328
9.5809 | -90.9339
9.5812 | -90.9351
9.5816 | -90.936
9.5819 | -90.9374
9.5823 | -90.9388
9.5827 | -90.9395
9.583 | -90.9404
9.5834 | -90.9415
9.5837 | -90.9429
9.5841 | -90.944
9.5845 | -90.9461
9.5848 | -90.9481
9.5852 | -90.9497
9.5856 | -90.9512
9.5859 | -90.952
9.5863 | -90.9525
9.5866 | -90.9532
9.587 | -90.954
9.5874 | -90.9551
9.5877 | -90.9556
9.5881 | -90.9563
9.5885 | -90.9571
9.5888 | -90.958
9.5892 | -90.9596
9.5895 | -90.961
9.5899 | -90.9623
9.5903 | -90.9638
9.5906 | -90.9646
9.591 | -90.9659
9.5913 | -90.967
9.5917 | -90.9681
9.5921 | -90.9691
9.5924 | -90.9703
9.5928 | -90.9716
9.5932 | -90.9726
9.5935 | -90.9751
9.5939 | -90.9768
9.5942 | -90.9778
9.5946 | -90.9795
9.595 | -90.9813
9.5953 | -90.9821
9.5957 | -90.9826
9.596 | -90.9832
9.5964 | -90.9838
9.5968 | -90.9846
9.5971 | -90.9856
9.5975 | -90.9865
9.5979 | -90.9872
9.5982 | -90.9881
9.5986 | -90.9893
9.5989 | -90.9908
9.5993 | -90.9919
9.5997 | -90.9928
9.6 | -90.9937
9.6004 | -90.9946
9.6008 | -90.9956
9.6011 | -90.9963
9.6015 | -90.9975
9.6018 | -90.9985
9.6022 | -91.0001
9.6026 | -91.0018
9.6029 | -91.0031
9.6033 | -91.0035
9.6036 | -91.0023
9.604 | -91.0032
9.6044 | -91.0041
9.6047 | -91.005
9.6051 | -91.006
9.6055 | -91.0064
9.6058 | -91.0074
9.6062 | -91.0084
9.6065 | -91.0094
9.6069 | -91.0103
9.6073 | -91.0112
9.6076 | -91.0122
9.608 | -91.013
9.6083 | -91.0136
9.6087 | -91.0146
9.6091 | -91.0152
9.6094 | -91.015
9.6098 | -91.0158
9.6102 | -91.017
9.6105 | -91.018
9.6109 | -91.0189
9.6112 | -91.0197
9.6116 | -91.0205
9.612 | -91.0214
9.6123 | -91.0223
9.6127 | -91.0232
9.6131 | -91.024
9.6134 | -91.0254
9.6138 | -91.0271
9.6141 | -91.028
9.6145 | -91.0288
9.6149 | -91.0295
9.6152 | -91.0312
9.6156 | -91.0328
9.6159 | -91.0337
9.6163 | -91.0354
9.6167 | -91.0368
9.617 | -91.0381
9.6174 | -91.0392
9.6178 | -91.0401
9.6181 | -91.0408
9.6185 | -91.0412
9.6188 | -91.0421
9.6192 | -91.043
9.6196 | -91.0438
9.6199 | -91.0449
9.6203 | -91.0456
9.6207 | -91.0463
9.621 | -91.0469
9.6214 | -91.0479
9.6217 | -91.049
9.6221 | -91.0501
9.6225 | -91.051
9.6228 | -91.0523
9.6232 | -91.0534
9.6235 | -91.0547
9.6239 | -91.057
9.6243 | -91.0576
9.6246 | -91.0596
9.625 | -91.0619
9.6254 | -91.0631
9.6257 | -91.0638
9.6261 | -91.0645
9.6264 | -91.0656
9.6268 | -91.0672
9.6272 | -91.0694
9.6275 | -91.0712
9.6279 | -91.0719
9.6282 | -91.0727
9.6286 | -91.0735
9.629 | -91.0743
9.6293 | -91.075
9.6297 | -91.076
9.6301 | -91.0766
9.6304 | -91.0773
9.6308 | -91.0782
9.6311 | -91.0791
9.6315 | -91.0802
9.6319 | -91.0814
9.6322 | -91.0827
9.6326 | -91.0842
9.633 | -91.0857
9.6333 | -91.0866
9.6337 | -91.0875
9.634 | -91.0883
9.6344 | -91.0892
9.6348 | -91.0897
9.6351 | -91.0904
9.6355 | -91.0914
9.6358 | -91.0923
9.6362 | -91.0932
9.6366 | -91.0939
9.6369 | -91.0938
9.6373 | -91.0944
9.6377 | -91.0955
9.638 | -91.0968
9.6384 | -91.0982
9.6387 | -91.0994
9.6391 | -91.101
9.6395 | -91.102
9.6398 | -91.1023
9.6402 | -91.1027
9.6406 | -91.1035
9.6409 | -91.1044
9.6413 | -91.1054
9.6416 | -91.1062
9.642 | -91.1072
9.6424 | -91.1085
9.6427 | -91.1093
9.6431 | -91.11
9.6434 | -91.111
9.6438 | -91.1111
9.6442 | -91.111
9.6445 | -91.1113
9.6449 | -91.1121
9.6453 | -91.1127
9.6456 | -91.1138
9.646 | -91.1148
9.6463 | -91.1155
9.6467 | -91.1165
9.6471 | -91.1184
9.6474 | -91.1209
9.6478 | -91.1224
9.6481 | -91.1232
9.6485 | -91.1239
9.6489 | -91.1246
9.6492 | -91.1255
9.6496 | -91.1271
9.65 | -91.1292
9.6503 | -91.1306
9.6507 | -91.1315
9.651 | -91.1327
9.6514 | -91.1336
9.6518 | -91.1346
9.6521 | -91.1358
9.6525 | -91.1365
9.6529 | -91.1373
9.6532 | -91.138
9.6536 | -91.1384
9.6539 | -91.1397
9.6543 | -91.1406
9.6547 | -91.1408
9.655 | -91.1414
9.6554 | -91.1424
9.6557 | -91.1425
9.6561 | -91.1438
9.6565 | -91.1448
9.6568 | -91.1455
9.6572 | -91.146
9.6576 | -91.1467
9.6579 | -91.1481
9.6583 | -91.1497
9.6586 | -91.1503
9.659 | -91.1514
9.6594 | -91.1521
9.6597 | -91.1528
9.6601 | -91.1537
9.6605 | -91.1542
9.6608 | -91.1551
9.6612 | -91.1562
9.6615 | -91.1571
9.6619 | -91.1574
9.6623 | -91.1583
9.6626 | -91.1591
9.663 | -91.1598
9.6633 | -91.1604
9.6637 | -91.1609
9.6641 | -91.1616
9.6644 | -91.1627
9.6648 | -91.1635
9.6652 | -91.1642
9.6655 | -91.1649
9.6659 | -91.1658
9.6662 | -91.1664
9.6666 | -91.1672
9.667 | -91.1675
9.6673 | -91.1681
9.6677 | -91.1689
9.668 | -91.1696
9.6684 | -91.1702
9.6688 | -91.1709
9.6691 | -91.1714
9.6695 | -91.1731
9.6699 | -91.1761
9.6702 | -91.1774
9.6706 | -91.1783
9.6709 | -91.1788
9.6713 | -91.1802
9.6717 | -91.1815
9.672 | -91.1817
9.6724 | -91.1827
9.6728 | -91.1839
9.6731 | -91.1851
9.6735 | -91.1857
9.6738 | -91.1865
9.6742 | -91.1871
9.6746 | -91.1873
9.6749 | -91.1879
9.6753 | -91.1886
9.6756 | -91.1894
9.676 | -91.1903
9.6764 | -91.1913
9.6767 | -91.1915
9.6771 | -91.1923
9.6775 | -91.193
9.6778 | -91.1942
9.6782 | -91.1953
9.6785 | -91.1958
9.6789 | -91.1956
9.6793 | -91.1956
9.6796 | -91.1958
9.68 | -91.1963
9.6803 | -91.1975
9.6807 | -91.1983
9.6811 | -91.199
9.6814 | -91.1995
9.6818 | -91.1996
9.6822 | -91.1997
9.6825 | -91.1999
9.6829 | -91.2003
9.6832 | -91.2014
9.6836 | -91.202
9.684 | -91.2024
9.6843 | -91.203
9.6847 | -91.2038
9.6851 | -91.2046
9.6854 | -91.2047
9.6858 | -91.2052
9.6861 | -91.206
9.6865 | -91.207
9.6869 | -91.2077
9.6872 | -91.2081
9.6876 | -91.2086
9.6879 | -91.2091
9.6883 | -91.2094
9.6887 | -91.2098
9.689 | -91.2101
9.6894 | -91.2105
9.6898 | -91.2111
9.6901 | -91.2117
9.6905 | -91.2122
9.6908 | -91.213
9.6912 | -91.214
9.6916 | -91.215
9.6919 | -91.2155
9.6923 | -91.2159
9.6927 | -91.2163
9.693 | -91.2171
9.6934 | -91.2177
9.6937 | -91.2184
9.6941 | -91.2192
9.6945 | -91.2201
9.6948 | -91.2208
9.6952 | -91.2212
9.6955 | -91.2216
9.6959 | -91.222
9.6963 | -91.2229
9.6966 | -91.2236
9.697 | -91.2243
9.6974 | -91.2252
9.6977 | -91.226
9.6981 | -91.2269
9.6984 | -91.2278
9.6988 | -91.2288
9.6992 | -91.2297
9.6995 | -91.2311
9.6999 | -91.2317
9.7002 | -91.2323
9.7006 | -91.2329
9.701 | -91.2333
9.7013 | -91.2339
9.7017 | -91.2352
9.7021 | -91.236
9.7024 | -91.2369
9.7028 | -91.2372
9.7031 | -91.2372
9.7035 | -91.2378
9.7039 | -91.2385
9.7042 | -91.2387
9.7046 | -91.2393
9.705 | -91.2396
9.7053 | -91.2405
9.7057 | -91.2416
9.706 | -91.2423
9.7064 | -91.2431
9.7068 | -91.2438
9.7071 | -91.2445
9.7075 | -91.2451
9.7078 | -91.2452
9.7082 | -91.2464
9.7086 | -91.2484
9.7089 | -91.2496
9.7093 | -91.2503
9.7097 | -91.251
9.71 | -91.2524
9.7104 | -91.2543
9.7107 | -91.2551
9.7111 | -91.2558
9.7115 | -91.2565
9.7118 | -91.2569
9.7122 | -91.2573
9.7125 | -91.2577
9.7129 | -91.2579
9.7133 | -91.2578
9.7136 | -91.259
9.714 | -91.2605
9.7144 | -91.2616
9.7147 | -91.2623
9.7151 | -91.2628
9.7154 | -91.2634
9.7158 | -91.2643
9.7162 | -91.265
9.7165 | -91.2656
9.7169 | -91.2667
9.7173 | -91.2674
9.7176 | -91.2678
9.718 | -91.2686
9.7183 | -91.2694
9.7187 | -91.2704
9.7191 | -91.2708
9.7194 | -91.2712
9.7198 | -91.272
9.7201 | -91.2728
9.7205 | -91.2737
9.7209 | -91.2743
9.7212 | -91.2746
9.7216 | -91.2749
9.722 | -91.2754
9.7223 | -91.2761
9.7227 | -91.2769
9.723 | -91.2775
9.7234 | -91.2783
9.7238 | -91.2791
9.7241 | -91.2795
9.7245 | -91.2799
9.7248 | -91.2803
9.7252 | -91.2807
9.7256 | -91.2815
9.7259 | -91.2822
9.7263 | -91.2832
9.7267 | -91.2841
9.727 | -91.2843
9.7274 | -91.2843
9.7277 | -91.2845
9.7281 | -91.2846
9.7285 | -91.2854
9.7288 | -91.2853
9.7292 | -91.2863
9.7296 | -91.2874
9.7299 | -91.288
9.7303 | -91.2882
9.7306 | -91.2884
9.731 | -91.2892
9.7314 | -91.2898
9.7317 | -91.2905
9.7321 | -91.2913
9.7324 | -91.2919
9.7328 | -91.2925
9.7332 | -91.2932
9.7335 | -91.2936
9.7339 | -91.2941
9.7343 | -91.295
9.7346 | -91.2958
9.735 | -91.2966
9.7353 | -91.2974
9.7357 | -91.2981
9.7361 | -91.2989
9.7364 | -91.2993
9.7368 | -91.3002
9.7372 | -91.3016
9.7375 | -91.3032
9.7379 | -91.3041
9.7382 | -91.3048
9.7386 | -91.3056
9.739 | -91.3064
9.7393 | -91.3071
9.7397 | -91.308
9.74 | -91.309
9.7404 | -91.3097
9.7408 | -91.3106
9.7411 | -91.312
9.7415 | -91.3134
9.7419 | -91.316
9.7422 | -91.3176
9.7426 | -91.3182
9.7429 | -91.3191
9.7433 | -91.3197
9.7437 | -91.3201
9.744 | -91.3213
9.7444 | -91.3223
9.7447 | -91.3232
9.7451 | -91.3239
9.7455 | -91.3244
9.7458 | -91.325
9.7462 | -91.3263
9.7466 | -91.3276
9.7469 | -91.3281
9.7473 | -91.3287
9.7476 | -91.3295
9.748 | -91.33
9.7484 | -91.3306
9.7487 | -91.3313
9.7491 | -91.3328
9.7495 | -91.3338
9.7498 | -91.3348
9.7502 | -91.3353
9.7505 | -91.3357
9.7509 | -91.3359
9.7513 | -91.3366
9.7516 | -91.3371
9.752 | -91.3372
9.7523 | -91.3376
9.7527 | -91.3391
9.7531 | -91.34
9.7534 | -91.3398
9.7538 | -91.3398
9.7542 | -91.3407
9.7545 | -91.3416
9.7549 | -91.3422
9.7552 | -91.3427
9.7556 | -91.3429
9.756 | -91.3432
9.7563 | -91.3443
9.7567 | -91.3453
9.7571 | -91.346
9.7574 | -91.3471
9.7578 | -91.3472
9.7581 | -91.3477
9.7585 | -91.3485
9.7589 | -91.3492
9.7592 | -91.3498
9.7596 | -91.3503
9.7599 | -91.351
9.7603 | -91.3512
9.7607 | -91.3518
9.761 | -91.3533
9.7614 | -91.3542
9.7618 | -91.3548
9.7621 | -91.356
9.7625 | -91.3563
9.7628 | -91.3569
9.7632 | -91.3576
9.7636 | -91.3578
9.7639 | -91.3582
9.7643 | -91.3584
9.7646 | -91.3587
9.765 | -91.359
9.7654 | -91.3593
9.7657 | -91.36
9.7661 | -91.3602
9.7665 | -91.3607
9.7668 | -91.3627
9.7672 | -91.3656
9.7675 | -91.3665
9.7679 | -91.3667
9.7683 | -91.3677
9.7686 | -91.3685
9.769 | -91.3693
9.7694 | -91.3699
9.7697 | -91.3704
9.7701 | -91.3707
9.7704 | -91.3715
9.7708 | -91.3708
9.7712 | -91.3714
9.7715 | -91.373
9.7719 | -91.3738
9.7722 | -91.3745
9.7726 | -91.3754
9.773 | -91.3762
9.7733 | -91.3773
9.7737 | -91.3791
9.7741 | -91.3804
9.7744 | -91.3812
9.7748 | -91.382
9.7751 | -91.3831
9.7755 | -91.3831
9.7759 | -91.3833
9.7762 | -91.3842
9.7766 | -91.3834
9.7769 | -91.382
9.7773 | -91.3818
9.7777 | -91.3823
9.778 | -91.3826
9.7784 | -91.3829
9.7788 | -91.3834
9.7791 | -91.3839
9.7795 | -91.3845
9.7798 | -91.3849
9.7802 | -91.3856
9.7806 | -91.3862
9.7809 | -91.3866
9.7813 | -91.3871
9.7817 | -91.3877
9.782 | -91.3882
9.7824 | -91.3888
9.7827 | -91.3898
9.7831 | -91.3906
9.7835 | -91.3913
9.7838 | -91.392
9.7842 | -91.3926
9.7845 | -91.393
9.7849 | -91.3929
9.7853 | -91.3929
9.7856 | -91.393
9.786 | -91.3928
9.7864 | -91.3922
9.7867 | -91.3912
9.7871 | -91.3912
9.7874 | -91.3918
9.7878 | -91.3924
9.7882 | -91.3931
9.7885 | -91.3946
9.7889 | -91.3953
9.7893 | -91.3957
9.7896 | -91.3962
9.79 | -91.3966
9.7903 | -91.3975
9.7907 | -91.3989
9.7911 | -91.3995
9.7914 | -91.4006
9.7918 | -91.4014
9.7921 | -91.4032
9.7925 | -91.4047
9.7929 | -91.4064
9.7932 | -91.4087
9.7936 | -91.4105
9.794 | -91.4111
9.7943 | -91.4112
9.7947 | -91.4116
9.795 | -91.412
9.7954 | -91.4128
9.7958 | -91.4134
9.7961 | -91.4136
9.7965 | -91.4145
9.7968 | -91.4158
9.7972 | -91.4166
9.7976 | -91.4176
9.7979 | -91.4186
9.7983 | -91.4199
9.7987 | -91.4205
9.799 | -91.421
9.7994 | -91.4214
9.7997 | -91.4227
9.8001 | -91.4237
9.8005 | -91.4245
9.8008 | -91.4249
9.8012 | -91.4245
9.8016 | -91.4246
9.8019 | -91.4252
9.8023 | -91.4257
9.8026 | -91.4264
9.803 | -91.4268
9.8034 | -91.4273
9.8037 | -91.4277
9.8041 | -91.4283
9.8044 | -91.4294
9.8048 | -91.4309
9.8052 | -91.4318
9.8055 | -91.4323
9.8059 | -91.4329
9.8063 | -91.4335
9.8066 | -91.4341
9.807 | -91.4347
9.8073 | -91.4364
9.8077 | -91.438
9.8081 | -91.4389
9.8084 | -91.4394
9.8088 | -91.4396
9.8092 | -91.4397
9.8095 | -91.4399
9.8099 | -91.4403
9.8102 | -91.4408
9.8106 | -91.4413
9.811 | -91.4418
9.8113 | -91.4423
9.8117 | -91.4427
9.812 | -91.4437
9.8124 | -91.4448
9.8128 | -91.4469
9.8131 | -91.4477
9.8135 | -91.4483
9.8139 | -91.4488
9.8142 | -91.4491
9.8146 | -91.45
9.8149 | -91.4513
9.8153 | -91.4531
9.8157 | -91.4551
9.816 | -91.4565
9.8164 | -91.4575
9.8167 | -91.4583
9.8171 | -91.4594
9.8175 | -91.4603
9.8178 | -91.4604
9.8182 | -91.4603
9.8186 | -91.4603
9.8189 | -91.4604
9.8193 | -91.4606
9.8196 | -91.4611
9.82 | -91.462
9.8204 | -91.4623
9.8207 | -91.4627
9.8211 | -91.4635
9.8215 | -91.4644
9.8218 | -91.4649
9.8222 | -91.466
9.8225 | -91.4669
9.8229 | -91.4675
9.8233 | -91.4681
9.8236 | -91.4686
9.824 | -91.4698
9.8243 | -91.471
9.8247 | -91.4715
9.8251 | -91.4718
9.8254 | -91.4724
9.8258 | -91.4735
9.8262 | -91.4743
9.8265 | -91.4746
9.8269 | -91.4749
9.8272 | -91.4754
9.8276 | -91.4758
9.828 | -91.4765
9.8283 | -91.477
9.8287 | -91.4773
9.829 | -91.4779
9.8294 | -91.478
9.8298 | -91.4789
9.8301 | -91.4798
9.8305 | -91.4804
9.8309 | -91.481
9.8312 | -91.4814
9.8316 | -91.4821
9.8319 | -91.4823
9.8323 | -91.4829
9.8327 | -91.484
9.833 | -91.4844
9.8334 | -91.4855
9.8338 | -91.4861
9.8341 | -91.4866
9.8345 | -91.4871
9.8348 | -91.4877
9.8352 | -91.4884
9.8356 | -91.4885
9.8359 | -91.4888
9.8363 | -91.4898
9.8366 | -91.4904
9.837 | -91.4909
9.8374 | -91.4911
9.8377 | -91.4906
9.8381 | -91.4907
9.8385 | -91.491
9.8388 | -91.4912
9.8392 | -91.4914
9.8395 | -91.492
9.8399 | -91.4928
9.8403 | -91.4936
9.8406 | -91.4946
9.841 | -91.4954
9.8414 | -91.4962
9.8417 | -91.4974
9.8421 | -91.4984
9.8424 | -91.499
9.8428 | -91.4993
9.8432 | -91.4999
9.8435 | -91.5008
9.8439 | -91.5014
9.8442 | -91.5019
9.8446 | -91.5023
9.845 | -91.5028
9.8453 | -91.503
9.8457 | -91.5032
9.8461 | -91.5034
9.8464 | -91.5034
9.8468 | -91.5038
9.8471 | -91.5044
9.8475 | -91.5051
9.8479 | -91.5056
9.8482 | -91.5059
9.8486 | -91.5063
9.8489 | -91.5072
9.8493 | -91.5077
9.8497 | -91.5086
9.85 | -91.5097
9.8504 | -91.5105
9.8508 | -91.5105
9.8511 | -91.5111
9.8515 | -91.5113
9.8518 | -91.512
9.8522 | -91.5128
9.8526 | -91.5138
9.8529 | -91.5149
9.8533 | -91.5153
9.8537 | -91.5156
9.854 | -91.516
9.8544 | -91.5164
9.8547 | -91.5168
9.8551 | -91.5171
9.8555 | -91.5177
9.8558 | -91.5182
9.8562 | -91.5188
9.8565 | -91.5193
9.8569 | -91.5202
9.8573 | -91.521
9.8576 | -91.5217
9.858 | -91.522
9.8584 | -91.5228
9.8587 | -91.5237
9.8591 | -91.5241
9.8594 | -91.5245
9.8598 | -91.5253
9.8602 | -91.5264
9.8605 | -91.5271
9.8609 | -91.5272
9.8612 | -91.5279
9.8616 | -91.5287
9.862 | -91.5295
9.8623 | -91.5302
9.8627 | -91.5307
9.8631 | -91.531
9.8634 | -91.5313
9.8638 | -91.5317
9.8641 | -91.5327
9.8645 | -91.5334
9.8649 | -91.5335
9.8652 | -91.5344
9.8656 | -91.5362
9.866 | -91.5377
9.8663 | -91.5386
9.8667 | -91.5395
9.867 | -91.5399
9.8674 | -91.5405
9.8678 | -91.5413
9.8681 | -91.5421
9.8685 | -91.5429
9.8688 | -91.5436
9.8692 | -91.5447
9.8696 | -91.5453
9.8699 | -91.5467
9.8703 | -91.5486
9.8707 | -91.5492
9.871 | -91.5498
9.8714 | -91.5505
9.8717 | -91.551
9.8721 | -91.5512
9.8725 | -91.5515
9.8728 | -91.552
9.8732 | -91.5529
9.8736 | -91.5541
9.8739 | -91.5551
9.8743 | -91.5558
9.8746 | -91.5561
9.875 | -91.5561
9.8754 | -91.5562
9.8757 | -91.5561
9.8761 | -91.5561
9.8764 | -91.5563
9.8768 | -91.5568
9.8772 | -91.5574
9.8775 | -91.5581
9.8779 | -91.5591
9.8783 | -91.5593
9.8786 | -91.5579
9.879 | -91.5579
9.8793 | -91.5584
9.8797 | -91.5587
9.8801 | -91.5595
9.8804 | -91.5603
9.8808 | -91.5609
9.8811 | -91.5618
9.8815 | -91.5625
9.8819 | -91.5629
9.8822 | -91.5638
9.8826 | -91.5631
9.883 | -91.562
9.8833 | -91.5623
9.8837 | -91.5627
9.884 | -91.5631
9.8844 | -91.5635
9.8848 | -91.5638
9.8851 | -91.5639
9.8855 | -91.5645
9.8859 | -91.5645
9.8862 | -91.5651
9.8866 | -91.566
9.8869 | -91.5671
9.8873 | -91.5682
9.8877 | -91.5688
9.888 | -91.5693
9.8884 | -91.5695
9.8887 | -91.5699
9.8891 | -91.5702
9.8895 | -91.5698
9.8898 | -91.5702
9.8902 | -91.5708
9.8906 | -91.5711
9.8909 | -91.5714
9.8913 | -91.5717
9.8916 | -91.5723
9.892 | -91.5736
9.8924 | -91.5743
9.8927 | -91.5749
9.8931 | -91.5752
9.8934 | -91.5758
9.8938 | -91.5763
9.8942 | -91.5768
9.8945 | -91.5771
9.8949 | -91.5775
9.8953 | -91.578
9.8956 | -91.5784
9.896 | -91.5776
9.8963 | -91.5781
9.8967 | -91.5785
9.8971 | -91.5791
9.8974 | -91.5796
9.8978 | -91.5804
9.8982 | -91.5811
9.8985 | -91.5815
9.8989 | -91.5819
9.8992 | -91.5826
9.8996 | -91.5831
9.9 | -91.5832
9.9003 | -91.5834
9.9007 | -91.5836
9.901 | -91.5836
9.9014 | -91.584
9.9018 | -91.5844
9.9021 | -91.5845
9.9025 | -91.5848
9.9029 | -91.5855
9.9032 | -91.587
9.9036 | -91.5884
9.9039 | -91.5892
9.9043 | -91.5891
9.9047 | -91.5887
9.905 | -91.589
9.9054 | -91.5899
9.9058 | -91.591
9.9061 | -91.5917
9.9065 | -91.5921
9.9068 | -91.5924
9.9072 | -91.5929
9.9076 | -91.5934
9.9079 | -91.5935
9.9083 | -91.5937
9.9086 | -91.5939
9.909 | -91.5949
9.9094 | -91.5964
9.9097 | -91.5973
9.9101 | -91.5977
9.9105 | -91.5981
9.9108 | -91.5987
9.9112 | -91.5988
9.9115 | -91.599
9.9119 | -91.5996
9.9123 | -91.5999
9.9126 | -91.6004
9.913 | -91.601
9.9133 | -91.602
9.9137 | -91.603
9.9141 | -91.6034
9.9144 | -91.6039
9.9148 | -91.6039
9.9152 | -91.6043
9.9155 | -91.6047
9.9159 | -91.605
9.9162 | -91.6051
9.9166 | -91.6057
9.917 | -91.6056
9.9173 | -91.606
9.9177 | -91.6065
9.9181 | -91.607
9.9184 | -91.6074
9.9188 | -91.6076
9.9191 | -91.6083
9.9195 | -91.6095
9.9199 | -91.6103
9.9202 | -91.6107
9.9206 | -91.6113
9.9209 | -91.6115
9.9213 | -91.6117
9.9217 | -91.6122
9.922 | -91.6125
9.9224 | -91.6125
9.9228 | -91.6127
9.9231 | -91.6129
9.9235 | -91.613
9.9238 | -91.6129
9.9242 | -91.6128
9.9246 | -91.6129
9.9249 | -91.6136
9.9253 | -91.6143
9.9256 | -91.6148
9.926 | -91.6152
9.9264 | -91.6153
9.9267 | -91.6153
9.9271 | -91.6159
9.9275 | -91.616
9.9278 | -91.6162
9.9282 | -91.6162
9.9285 | -91.6161
9.9289 | -91.616
9.9293 | -91.6157
9.9296 | -91.6159
9.93 | -91.6167
9.9304 | -91.6173
9.9307 | -91.6179
9.9311 | -91.6186
9.9314 | -91.6197
9.9318 | -91.6207
9.9322 | -91.6213
9.9325 | -91.6219
9.9329 | -91.6225
9.9332 | -91.6233
9.9336 | -91.6239
9.934 | -91.6244
9.9343 | -91.6249
9.9347 | -91.6253
9.9351 | -91.626
9.9354 | -91.6267
9.9358 | -91.6272
9.9361 | -91.6278
9.9365 | -91.6283
9.9369 | -91.6287
9.9372 | -91.6291
9.9376 | -91.6297
9.938 | -91.6301
9.9383 | -91.6303
9.9387 | -91.6304
9.939 | -91.6308
9.9394 | -91.6311
9.9398 | -91.6317
9.9401 | -91.6326
9.9405 | -91.6338
9.9408 | -91.6342
9.9412 | -91.6345
9.9416 | -91.6351
9.9419 | -91.6354
9.9423 | -91.6357
9.9427 | -91.6363
9.943 | -91.6372
9.9434 | -91.6379
9.9437 | -91.6386
9.9441 | -91.6394
9.9445 | -91.6405
9.9448 | -91.641
9.9452 | -91.6408
9.9455 | -91.6405
9.9459 | -91.6408
9.9463 | -91.6414
9.9466 | -91.6419
9.947 | -91.6429
9.9474 | -91.6439
9.9477 | -91.6445
9.9481 | -91.6448
9.9484 | -91.6456
9.9488 | -91.6464
9.9492 | -91.6471
9.9495 | -91.6473
9.9499 | -91.6479
9.9503 | -91.6486
9.9506 | -91.649
9.951 | -91.6494
9.9513 | -91.6497
9.9517 | -91.6499
9.9521 | -91.6503
9.9524 | -91.6508
9.9528 | -91.6509
9.9531 | -91.6509
9.9535 | -91.6512
9.9539 | -91.6516
9.9542 | -91.652
9.9546 | -91.6525
9.955 | -91.6527
9.9553 | -91.6521
9.9557 | -91.6507
9.956 | -91.65
9.9564 | -91.65
9.9568 | -91.6503
9.9571 | -91.6506
9.9575 | -91.6512
9.9579 | -91.6517
9.9582 | -91.652
9.9586 | -91.6522
9.9589 | -91.6524
9.9593 | -91.6527
9.9597 | -91.6523
9.96 | -91.653
9.9604 | -91.6539
9.9607 | -91.6542
9.9611 | -91.6546
9.9615 | -91.6547
9.9618 | -91.6548
9.9622 | -91.6552
9.9626 | -91.6564
9.9629 | -91.6566
9.9633 | -91.657
9.9636 | -91.6574
9.964 | -91.658
9.9644 | -91.6583
9.9647 | -91.6586
9.9651 | -91.6592
9.9654 | -91.6596
9.9658 | -91.6596
9.9662 | -91.6595
9.9665 | -91.6598
9.9669 | -91.6602
9.9673 | -91.6605
9.9676 | -91.661
9.968 | -91.6617
9.9683 | -91.6627
9.9687 | -91.6638
9.9691 | -91.6644
9.9694 | -91.665
9.9698 | -91.6653
9.9702 | -91.6657
9.9705 | -91.6661
9.9709 | -91.6664
9.9712 | -91.667
9.9716 | -91.6673
9.972 | -91.6676
9.9723 | -91.6682
9.9727 | -91.6687
9.973 | -91.6689
9.9734 | -91.6691
9.9738 | -91.6694
9.9741 | -91.6709
9.9745 | -91.6723
9.9749 | -91.6725
9.9752 | -91.6728
9.9756 | -91.6727
9.9759 | -91.6733
9.9763 | -91.6744
9.9767 | -91.6748
9.977 | -91.6746
9.9774 | -91.6742
9.9777 | -91.6746
9.9781 | -91.6752
9.9785 | -91.6754
9.9788 | -91.6752
9.9792 | -91.6748
9.9796 | -91.6753
9.9799 | -91.6754
9.9803 | -91.6758
9.9806 | -91.6763
9.981 | -91.6767
9.9814 | -91.6771
9.9817 | -91.6776
9.9821 | -91.678
9.9825 | -91.6782
9.9828 | -91.6787
9.9832 | -91.6799
9.9835 | -91.6823
9.9839 | -91.6829
9.9843 | -91.6829
9.9846 | -91.6833
9.985 | -91.6835
9.9853 | -91.6837
9.9857 | -91.6839
9.9861 | -91.6844
9.9864 | -91.6846
9.9868 | -91.6847
9.9872 | -91.685
9.9875 | -91.6853
9.9879 | -91.6855
9.9882 | -91.6855
9.9886 | -91.6862
9.989 | -91.6864
9.9893 | -91.6868
9.9897 | -91.6866
9.9901 | -91.6866
9.9904 | -91.6872
9.9908 | -91.6876
9.9911 | -91.6884
9.9915 | -91.6895
9.9919 | -91.6907
9.9922 | -91.6912
9.9926 | -91.6912
9.9929 | -91.6915
9.9933 | -91.692
9.9937 | -91.6922
9.994 | -91.6927
9.9944 | -91.6931
9.9948 | -91.6932
9.9951 | -91.6935
9.9955 | -91.6936
9.9958 | -91.6935
9.9962 | -91.6934
9.9966 | -91.6941
9.9969 | -91.6952
9.9973 | -91.6958
9.9976 | -91.6966
9.998 | -91.697
9.9984 | -91.6974
9.9987 | -91.6978
9.9991 | -91.6985
9.9995 | -91.6989
9.9998 | -91.6992
10.0002 | -91.6997
10.0005 | -91.7
10.0009 | -91.7003
10.0013 | -91.7003
10.0016 | -91.6996
10.002 | -91.6986
10.0024 | -91.6977
10.0027 | -91.6978
10.0031 | -91.698
10.0034 | -91.6982
10.0038 | -91.6985
10.0042 | -91.6994
10.0045 | -91.7
10.0049 | -91.6999
10.0052 | -91.6997
10.0056 | -91.7003
10.006 | -91.7009
10.0063 | -91.7013
10.0067 | -91.7017
10.0071 | -91.7021
10.0074 | -91.7023
10.0078 | -91.702
10.0081 | -91.701
10.0085 | -91.6999
10.0089 | -91.6999
10.0092 | -91.7006
10.0096 | -91.7013
10.0099 | -91.702
10.0103 | -91.7028
10.0107 | -91.7025
10.011 | -91.7022
10.0114 | -91.7026
10.0118 | -91.7038
10.0121 | -91.7048
10.0125 | -91.7049
10.0128 | -91.7048
10.0132 | -91.7049
10.0136 | -91.7054
10.0139 | -91.7061
10.0143 | -91.706
10.0147 | -91.706
10.015 | -91.7061
10.0154 | -91.7064
10.0157 | -91.7065
10.0161 | -91.7067
10.0165 | -91.7068
10.0168 | -91.707
10.0172 | -91.7071
10.0175 | -91.7074
10.0179 | -91.7077
10.0183 | -91.7081
10.0186 | -91.7082
10.019 | -91.708
10.0194 | -91.7083
10.0197 | -91.7089
10.0201 | -91.7096
10.0204 | -91.7099
10.0208 | -91.7099
10.0212 | -91.7101
10.0215 | -91.7098
10.0219 | -91.7095
10.0223 | -91.7096
10.0226 | -91.7099
10.023 | -91.7104
10.0233 | -91.7105
10.0237 | -91.7105
10.0241 | -91.7109
10.0244 | -91.7114
10.0248 | -91.7123
10.0251 | -91.713
10.0255 | -91.7132
10.0259 | -91.7135
10.0262 | -91.7137
10.0266 | -91.7133
10.027 | -91.7133
10.0273 | -91.7135
10.0277 | -91.7136
10.028 | -91.7141
10.0284 | -91.7146
10.0288 | -91.7147
10.0291 | -91.715
10.0295 | -91.7152
10.0298 | -91.7149
10.0302 | -91.7144
10.0306 | -91.7144
10.0309 | -91.7145
10.0313 | -91.7153
10.0317 | -91.7155
10.032 | -91.7162
10.0324 | -91.7161
10.0327 | -91.7162
10.0331 | -91.7165
10.0335 | -91.7166
10.0338 | -91.7173
10.0342 | -91.7184
10.0346 | -91.7188
10.0349 | -91.719
10.0353 | -91.7192
10.0356 | -91.7195
10.036 | -91.7198
10.0364 | -91.7201
10.0367 | -91.7203
10.0371 | -91.7204
10.0374 | -91.7205
10.0378 | -91.721
10.0382 | -91.7216
10.0385 | -91.7229
10.0389 | -91.7238
10.0393 | -91.7241
10.0396 | -91.7244
10.04 | -91.7247
10.0403 | -91.7247
10.0407 | -91.7257
10.0411 | -91.7278
10.0414 | -91.7291
10.0418 | -91.7299
10.0421 | -91.7304
10.0425 | -91.7307
10.0429 | -91.7311
10.0432 | -91.7303
10.0436 | -91.7301
10.044 | -91.7308
10.0443 | -91.7306
10.0447 | -91.7297
10.045 | -91.7287
10.0454 | -91.7286
10.0458 | -91.7289
10.0461 | -91.7293
10.0465 | -91.7298
10.0469 | -91.7303
10.0472 | -91.7304
10.0476 | -91.7305
10.0479 | -91.7311
10.0483 | -91.7318
10.0487 | -91.7322
10.049 | -91.7324
10.0494 | -91.7328
10.0497 | -91.733
10.0501 | -91.7328
10.0505 | -91.7325
10.0508 | -91.7326
10.0512 | -91.7331
10.0516 | -91.7331
10.0519 | -91.732
10.0523 | -91.7314
10.0526 | -91.7313
10.053 | -91.7311
10.0534 | -91.731
10.0537 | -91.7309
10.0541 | -91.7306
10.0545 | -91.7306
10.0548 | -91.7304
10.0552 | -91.7313
10.0555 | -91.732
10.0559 | -91.732
10.0563 | -91.7322
10.0566 | -91.7321
10.057 | -91.7317
10.0573 | -91.7312
10.0577 | -91.7313
10.0581 | -91.7318
10.0584 | -91.7313
10.0588 | -91.7305
10.0592 | -91.7303
10.0595 | -91.7307
10.0599 | -91.7307
10.0602 | -91.7308
10.0606 | -91.7309
10.061 | -91.7311
10.0613 | -91.7314
10.0617 | -91.7313
10.062 | -91.7309
10.0624 | -91.7306
10.0628 | -91.731
10.0631 | -91.7312
10.0635 | -91.7312
10.0639 | -91.7313
10.0642 | -91.7315
10.0646 | -91.7316
10.0649 | -91.7313
10.0653 | -91.7309
10.0657 | -91.7308
10.066 | -91.7308
10.0664 | -91.7308
10.0668 | -91.7304
10.0671 | -91.7301
10.0675 | -91.73
10.0678 | -91.7301
10.0682 | -91.7306
10.0686 | -91.731
10.0689 | -91.7313
10.0693 | -91.7317
10.0696 | -91.7323
10.07 | -91.7329
10.0704 | -91.733
10.0707 | -91.7331
10.0711 | -91.7334
10.0715 | -91.7336
10.0718 | -91.734
10.0722 | -91.7347
10.0725 | -91.7351
10.0729 | -91.7351
10.0733 | -91.735
10.0736 | -91.7351
10.074 | -91.735
10.0744 | -91.7347
10.0747 | -91.7351
10.0751 | -91.7353
10.0754 | -91.7353
10.0758 | -91.7354
10.0762 | -91.7358
10.0765 | -91.736
10.0769 | -91.7363
10.0772 | -91.7364
10.0776 | -91.7362
10.078 | -91.7358
10.0783 | -91.7352
10.0787 | -91.7353
10.0791 | -91.7363
10.0794 | -91.7367
10.0798 | -91.7368
10.0801 | -91.7376
10.0805 | -91.7383
10.0809 | -91.7386
10.0812 | -91.7385
10.0816 | -91.7384
10.0819 | -91.738
10.0823 | -91.7377
10.0827 | -91.7379
10.083 | -91.7377
10.0834 | -91.7377
10.0838 | -91.7378
10.0841 | -91.7372
10.0845 | -91.7372
10.0848 | -91.7375
10.0852 | -91.7375
10.0856 | -91.738
10.0859 | -91.7386
10.0863 | -91.7387
10.0867 | -91.7387
10.087 | -91.7385
10.0874 | -91.7385
10.0877 | -91.7385
10.0881 | -91.7392
10.0885 | -91.7398
10.0888 | -91.7396
10.0892 | -91.7399
10.0895 | -91.7406
10.0899 | -91.7409
10.0903 | -91.741
10.0906 | -91.7408
10.091 | -91.7409
10.0914 | -91.7409
10.0917 | -91.7402
10.0921 | -91.7399
10.0924 | -91.7399
10.0928 | -91.7402
10.0932 | -91.7402
10.0935 | -91.7404
10.0939 | -91.7405
10.0942 | -91.7403
10.0946 | -91.741
10.095 | -91.7415
10.0953 | -91.7423
10.0957 | -91.7426
10.0961 | -91.7423
10.0964 | -91.7427
10.0968 | -91.7434
10.0971 | -91.7435
10.0975 | -91.7431
10.0979 | -91.7431
10.0982 | -91.7431
10.0986 | -91.7443
10.099 | -91.746
10.0993 | -91.7466
10.0997 | -91.7466
10.1 | -91.7465
10.1004 | -91.7466
10.1008 | -91.7466
10.1011 | -91.7468
10.1015 | -91.7468
10.1018 | -91.7465
10.1022 | -91.7467
10.1026 | -91.7469
10.1029 | -91.747
10.1033 | -91.7475
10.1037 | -91.748
10.104 | -91.7479
10.1044 | -91.7475
10.1047 | -91.7469
10.1051 | -91.7478
10.1055 | -91.7495
10.1058 | -91.75
10.1062 | -91.7504
10.1066 | -91.7504
10.1069 | -91.7506
10.1073 | -91.7505
10.1076 | -91.7506
10.108 | -91.7509
10.1084 | -91.7509
10.1087 | -91.751
10.1091 | -91.751
10.1094 | -91.7507
10.1098 | -91.7506
10.1102 | -91.751
10.1105 | -91.7518
10.1109 | -91.7529
10.1113 | -91.7542
10.1116 | -91.7544
10.112 | -91.7542
10.1123 | -91.754
10.1127 | -91.754
10.1131 | -91.7541
10.1134 | -91.7545
10.1138 | -91.7548
10.1141 | -91.7552
10.1145 | -91.7558
10.1149 | -91.756
10.1152 | -91.7563
10.1156 | -91.7559
10.116 | -91.7555
10.1163 | -91.7554
10.1167 | -91.7552
10.117 | -91.7554
10.1174 | -91.7556
10.1178 | -91.7556
10.1181 | -91.7556
10.1185 | -91.7564
10.1189 | -91.7566
10.1192 | -91.7571
10.1196 | -91.7571
10.1199 | -91.7572
10.1203 | -91.7567
10.1207 | -91.7571
10.121 | -91.7575
10.1214 | -91.7572
10.1217 | -91.7571
10.1221 | -91.7573
10.1225 | -91.7577
10.1228 | -91.7583
10.1232 | -91.7582
10.1236 | -91.7582
10.1239 | -91.7585
10.1243 | -91.7587
10.1246 | -91.7589
10.125 | -91.7592
10.1254 | -91.7596
10.1257 | -91.7598
10.1261 | -91.76
10.1264 | -91.7601
10.1268 | -91.7605
10.1272 | -91.7609
10.1275 | -91.7605
10.1279 | -91.7604
10.1283 | -91.7605
10.1286 | -91.7612
10.129 | -91.7618
10.1293 | -91.7616
10.1297 | -91.7616
10.1301 | -91.7621
10.1304 | -91.763
10.1308 | -91.7635
10.1312 | -91.7642
10.1315 | -91.7643
10.1319 | -91.7643
10.1322 | -91.7644
10.1326 | -91.7639
10.133 | -91.7639
10.1333 | -91.7642
10.1337 | -91.7644
10.134 | -91.7648
10.1344 | -91.7655
10.1348 | -91.7656
10.1351 | -91.7658
10.1355 | -91.7662
10.1359 | -91.7666
10.1362 | -91.7668
10.1366 | -91.7669
10.1369 | -91.7668
10.1373 | -91.7665
10.1377 | -91.7663
10.138 | -91.7665
10.1384 | -91.7672
10.1388 | -91.768
10.1391 | -91.7683
10.1395 | -91.7689
10.1398 | -91.7706
10.1402 | -91.7721
10.1406 | -91.7731
10.1409 | -91.7737
10.1413 | -91.7737
10.1416 | -91.7741
10.142 | -91.7753
10.1424 | -91.7756
10.1427 | -91.7751
10.1431 | -91.7746
10.1435 | -91.7749
10.1438 | -91.7756
10.1442 | -91.7759
10.1445 | -91.7761
10.1449 | -91.7763
10.1453 | -91.7764
10.1456 | -91.7764
10.146 | -91.7767
10.1463 | -91.7773
10.1467 | -91.7776
10.1471 | -91.7775
10.1474 | -91.7775
10.1478 | -91.7777
10.1482 | -91.7778
10.1485 | -91.778
10.1489 | -91.7779
10.1492 | -91.7775
10.1496 | -91.7773
10.15 | -91.7775
10.1503 | -91.7776
10.1507 | -91.7778
10.1511 | -91.7778
10.1514 | -91.7777
10.1518 | -91.7775
10.1521 | -91.7772
10.1525 | -91.7782
10.1529 | -91.7795
10.1532 | -91.7799
10.1536 | -91.7799
10.1539 | -91.7801
10.1543 | -91.7802
10.1547 | -91.7796
10.155 | -91.7794
10.1554 | -91.7795
10.1558 | -91.7797
10.1561 | -91.7799
10.1565 | -91.7802
10.1568 | -91.781
10.1572 | -91.7812
10.1576 | -91.7815
10.1579 | -91.7819
10.1583 | -91.7821
10.1586 | -91.7819
10.159 | -91.7823
10.1594 | -91.7825
10.1597 | -91.7825
10.1601 | -91.783
10.1605 | -91.7835
10.1608 | -91.7834
10.1612 | -91.7834
10.1615 | -91.7836
10.1619 | -91.7842
10.1623 | -91.7843
10.1626 | -91.7842
10.163 | -91.7843
10.1634 | -91.7847
10.1637 | -91.7866
10.1641 | -91.788
10.1644 | -91.7881
10.1648 | -91.7882
10.1652 | -91.7881
10.1655 | -91.7882
10.1659 | -91.7883
10.1662 | -91.7887
10.1666 | -91.7892
10.167 | -91.7897
10.1673 | -91.7906
10.1677 | -91.7908
10.1681 | -91.791
10.1684 | -91.7912
10.1688 | -91.7916
10.1691 | -91.7919
10.1695 | -91.7924
10.1699 | -91.7928
10.1702 | -91.7928
10.1706 | -91.7921
10.171 | -91.7918
10.1713 | -91.7921
10.1717 | -91.7923
10.172 | -91.7925
10.1724 | -91.7928
10.1728 | -91.7928
10.1731 | -91.7927
10.1735 | -91.793
10.1738 | -91.7932
10.1742 | -91.7934
10.1746 | -91.7936
10.1749 | -91.794
10.1753 | -91.7948
10.1757 | -91.7944
10.176 | -91.7942
10.1764 | -91.7943
10.1767 | -91.7944
10.1771 | -91.7946
10.1775 | -91.795
10.1778 | -91.7959
10.1782 | -91.7962
10.1785 | -91.7963
10.1789 | -91.7968
10.1793 | -91.7978
10.1796 | -91.7986
10.18 | -91.799
10.1804 | -91.7993
10.1807 | -91.7995
10.1811 | -91.7997
10.1814 | -91.7999
10.1818 | -91.8003
10.1822 | -91.8007
10.1825 | -91.8015
10.1829 | -91.8021
10.1833 | -91.8024
10.1836 | -91.8026
10.184 | -91.8028
10.1843 | -91.8027
10.1847 | -91.8026
10.1851 | -91.8028
10.1854 | -91.8037
10.1858 | -91.804
10.1861 | -91.8038
10.1865 | -91.8041
10.1869 | -91.8043
10.1872 | -91.8047
10.1876 | -91.805
10.188 | -91.8047
10.1883 | -91.8052
10.1887 | -91.8067
10.189 | -91.8075
10.1894 | -91.8078
10.1898 | -91.8081
10.1901 | -91.8082
10.1905 | -91.8082
10.1908 | -91.8086
10.1912 | -91.8088
10.1916 | -91.8091
10.1919 | -91.8093
10.1923 | -91.8096
10.1927 | -91.8096
10.193 | -91.8096
10.1934 | -91.8099
10.1937 | -91.8102
10.1941 | -91.8107
10.1945 | -91.8116
10.1948 | -91.8124
10.1952 | -91.8127
10.1956 | -91.813
10.1959 | -91.8136
10.1963 | -91.8139
10.1966 | -91.8139
10.197 | -91.8143
10.1974 | -91.8148
10.1977 | -91.8148
10.1981 | -91.8144
10.1984 | -91.8136
10.1988 | -91.8134
10.1992 | -91.8134
10.1995 | -91.8135
10.1999 | -91.814
10.2003 | -91.8148
10.2006 | -91.8156
10.201 | -91.816
10.2013 | -91.8166
10.2017 | -91.8169
10.2021 | -91.8171
10.2024 | -91.8176
10.2028 | -91.8178
10.2032 | -91.8181
10.2035 | -91.8183
10.2039 | -91.8188
10.2042 | -91.8191
10.2046 | -91.8194
10.205 | -91.8201
10.2053 | -91.8205
10.2057 | -91.8206
10.206 | -91.8207
10.2064 | -91.8209
10.2068 | -91.8215
10.2071 | -91.8217
10.2075 | -91.8221
10.2079 | -91.823
10.2082 | -91.8237
10.2086 | -91.824
10.2089 | -91.8248
10.2093 | -91.8252
10.2097 | -91.8255
10.21 | -91.8259
10.2104 | -91.8255
10.2107 | -91.8259
10.2111 | -91.8263
10.2115 | -91.8265
10.2118 | -91.8267
10.2122 | -91.827
10.2126 | -91.8272
10.2129 | -91.828
10.2133 | -91.8289
10.2136 | -91.8298
10.214 | -91.8309
10.2144 | -91.8312
10.2147 | -91.8314
10.2151 | -91.8318
10.2155 | -91.8321
10.2158 | -91.8321
10.2162 | -91.8322
10.2165 | -91.833
10.2169 | -91.8328
10.2173 | -91.8331
10.2176 | -91.8331
10.218 | -91.8324
10.2183 | -91.8321
10.2187 | -91.8322
10.2191 | -91.8323
10.2194 | -91.833
10.2198 | -91.8336
10.2202 | -91.8339
10.2205 | -91.8342
10.2209 | -91.8346
10.2212 | -91.8345
10.2216 | -91.8347
10.222 | -91.8349
10.2223 | -91.8353
10.2227 | -91.8362
10.2231 | -91.8366
10.2234 | -91.8368
10.2238 | -91.8369
10.2241 | -91.8372
10.2245 | -91.8376
10.2249 | -91.838
10.2252 | -91.8388
10.2256 | -91.8392
10.2259 | -91.8393
10.2263 | -91.8399
10.2267 | -91.8404
10.227 | -91.8409
10.2274 | -91.8414
10.2278 | -91.8413
10.2281 | -91.8408
10.2285 | -91.8402
10.2288 | -91.8405
10.2292 | -91.841
10.2296 | -91.8418
10.2299 | -91.8425
10.2303 | -91.8432
10.2306 | -91.8441
10.231 | -91.8446
10.2314 | -91.845
10.2317 | -91.8457
10.2321 | -91.8459
10.2325 | -91.8461
10.2328 | -91.8468
10.2332 | -91.8472
10.2335 | -91.8478
10.2339 | -91.8482
10.2343 | -91.8486
10.2346 | -91.849
10.235 | -91.8498
10.2354 | -91.851
10.2357 | -91.8518
10.2361 | -91.852
10.2364 | -91.8524
10.2368 | -91.8523
10.2372 | -91.8523
10.2375 | -91.8522
10.2379 | -91.8529
10.2382 | -91.8539
10.2386 | -91.8542
10.239 | -91.8544
10.2393 | -91.8543
10.2397 | -91.8549
10.2401 | -91.8553
10.2404 | -91.8557
10.2408 | -91.8565
10.2411 | -91.8575
10.2415 | -91.8581
10.2419 | -91.8587
10.2422 | -91.8593
10.2426 | -91.8597
10.2429 | -91.8598
10.2433 | -91.8603
10.2437 | -91.8605
10.244 | -91.8607
10.2444 | -91.8604
10.2448 | -91.861
10.2451 | -91.8624
10.2455 | -91.8631
10.2458 | -91.8635
10.2462 | -91.8639
10.2466 | -91.8644
10.2469 | -91.8648
10.2473 | -91.8654
10.2477 | -91.866
10.248 | -91.867
10.2484 | -91.8682
10.2487 | -91.8686
10.2491 | -91.8693
10.2495 | -91.8696
10.2498 | -91.8695
10.2502 | -91.8703
10.2505 | -91.8707
10.2509 | -91.8715
10.2513 | -91.8726
10.2516 | -91.8741
10.252 | -91.8747
10.2524 | -91.8752
10.2527 | -91.876
10.2531 | -91.8766
10.2534 | -91.8776
10.2538 | -91.8783
10.2542 | -91.8785
10.2545 | -91.8789
10.2549 | -91.8793
10.2553 | -91.8796
10.2556 | -91.8786
10.256 | -91.8773
10.2563 | -91.8765
10.2567 | -91.8763
10.2571 | -91.877
10.2574 | -91.8775
10.2578 | -91.8778
10.2581 | -91.8778
10.2585 | -91.8781
10.2589 | -91.8785
10.2592 | -91.8792
10.2596 | -91.8801
10.26 | -91.8811
10.2603 | -91.8816
10.2607 | -91.8824
10.261 | -91.883
10.2614 | -91.8837
10.2618 | -91.8843
10.2621 | -91.8847
10.2625 | -91.8855
10.2628 | -91.8858
10.2632 | -91.8861
10.2636 | -91.8866
10.2639 | -91.8866
10.2643 | -91.8866
10.2647 | -91.8869
10.265 | -91.8876
10.2654 | -91.8882
10.2657 | -91.8886
10.2661 | -91.8883
10.2665 | -91.8894
10.2668 | -91.8905
10.2672 | -91.8904
10.2676 | -91.8911
10.2679 | -91.8914
10.2683 | -91.8916
10.2686 | -91.8921
10.269 | -91.892
10.2694 | -91.8923
10.2697 | -91.8928
10.2701 | -91.8932
10.2704 | -91.8937
10.2708 | -91.8943
10.2712 | -91.8951
10.2715 | -91.8953
10.2719 | -91.8958
10.2723 | -91.8964
10.2726 | -91.8972
10.273 | -91.898
10.2733 | -91.8983
10.2737 | -91.8983
10.2741 | -91.8987
10.2744 | -91.8992
10.2748 | -91.8997
10.2752 | -91.9005
10.2755 | -91.9016
10.2759 | -91.9023
10.2762 | -91.9029
10.2766 | -91.9036
10.277 | -91.9049
10.2773 | -91.9056
10.2777 | -91.906
10.278 | -91.9066
10.2784 | -91.9072
10.2788 | -91.9076
10.2791 | -91.908
10.2795 | -91.9085
10.2799 | -91.9093
10.2802 | -91.9098
10.2806 | -91.9102
10.2809 | -91.9107
10.2813 | -91.9112
10.2817 | -91.9108
10.282 | -91.911
10.2824 | -91.9111
10.2827 | -91.9114
10.2831 | -91.9116
10.2835 | -91.9121
10.2838 | -91.9125
10.2842 | -91.9131
10.2846 | -91.9134
10.2849 | -91.9138
10.2853 | -91.9144
10.2856 | -91.9153
10.286 | -91.9164
10.2864 | -91.9174
10.2867 | -91.9178
10.2871 | -91.9182
10.2875 | -91.9186
10.2878 | -91.9192
10.2882 | -91.9197
10.2885 | -91.9203
10.2889 | -91.9208
10.2893 | -91.9214
10.2896 | -91.9217
10.29 | -91.9218
10.2903 | -91.9224
10.2907 | -91.9234
10.2911 | -91.9241
10.2914 | -91.924
10.2918 | -91.9242
10.2922 | -91.9247
10.2925 | -91.9253
10.2929 | -91.926
10.2932 | -91.9263
10.2936 | -91.9273
10.294 | -91.9294
10.2943 | -91.931
10.2947 | -91.9314
10.295 | -91.9322
10.2954 | -91.9329
10.2958 | -91.933
10.2961 | -91.9331
10.2965 | -91.9334
10.2969 | -91.9338
10.2972 | -91.9344
10.2976 | -91.9349
10.2979 | -91.9352
10.2983 | -91.9356
10.2987 | -91.9353
10.299 | -91.9358
10.2994 | -91.9363
10.2998 | -91.9373
10.3001 | -91.9391
10.3005 | -91.94
10.3008 | -91.94
10.3012 | -91.9402
10.3016 | -91.9406
10.3019 | -91.9411
10.3023 | -91.9423
10.3026 | -91.9433
10.303 | -91.9442
10.3034 | -91.9446
10.3037 | -91.9452
10.3041 | -91.9457
10.3045 | -91.9461
10.3048 | -91.9464
10.3052 | -91.9468
10.3055 | -91.9473
10.3059 | -91.9477
10.3063 | -91.9483
10.3066 | -91.9483
10.307 | -91.9483
10.3073 | -91.9489
10.3077 | -91.95
10.3081 | -91.951
10.3084 | -91.951
10.3088 | -91.9506
10.3092 | -91.9506
10.3095 | -91.9505
10.3099 | -91.951
10.3102 | -91.9514
10.3106 | -91.9517
10.311 | -91.9522
10.3113 | -91.9525
10.3117 | -91.953
10.3121 | -91.9538
10.3124 | -91.9545
10.3128 | -91.9551
10.3131 | -91.9554
10.3135 | -91.9563
10.3139 | -91.957
10.3142 | -91.9575
10.3146 | -91.9579
10.3149 | -91.9584
10.3153 | -91.9587
10.3157 | -91.9588
10.316 | -91.9592
10.3164 | -91.9599
10.3168 | -91.9603
10.3171 | -91.9607
10.3175 | -91.9608
10.3178 | -91.9614
10.3182 | -91.9619
10.3186 | -91.9623
10.3189 | -91.9625
10.3193 | -91.9629
10.3197 | -91.9634
10.32 | -91.9643
10.3204 | -91.9648
10.3207 | -91.9654
10.3211 | -91.9662
10.3215 | -91.9674
10.3218 | -91.9685
10.3222 | -91.9695
10.3225 | -91.9701
10.3229 | -91.9707
10.3233 | -91.9713
10.3236 | -91.9713
10.324 | -91.97
10.3244 | -91.971
10.3247 | -91.9714
10.3251 | -91.9714
10.3254 | -91.9714
10.3258 | -91.9716
10.3262 | -91.972
10.3265 | -91.9726
10.3269 | -91.973
10.3272 | -91.9733
10.3276 | -91.974
10.328 | -91.9744
10.3283 | -91.9745
10.3287 | -91.9747
10.3291 | -91.9752
10.3294 | -91.9762
10.3298 | -91.9782
10.3301 | -91.979
10.3305 | -91.9796
10.3309 | -91.98
10.3312 | -91.9802
10.3316 | -91.9805
10.332 | -91.9809
10.3323 | -91.9815
10.3327 | -91.9819
10.333 | -91.9822
10.3334 | -91.9823
10.3338 | -91.9826
10.3341 | -91.9829
10.3345 | -91.9834
10.3348 | -91.984
10.3352 | -91.9844
10.3356 | -91.9848
10.3359 | -91.9852
10.3363 | -91.9856
10.3367 | -91.9858
10.337 | -91.9869
10.3374 | -91.9884
10.3377 | -91.9894
10.3381 | -91.9901
10.3385 | -91.9911
10.3388 | -91.9915
10.3392 | -91.9919
10.3395 | -91.9922
10.3399 | -91.9922
10.3403 | -91.9922
10.3406 | -91.9923
10.341 | -91.9925
10.3414 | -91.993
10.3417 | -91.9935
10.3421 | -91.9941
10.3424 | -91.9944
10.3428 | -91.995
10.3432 | -91.9953
10.3435 | -91.9958
10.3439 | -91.9962
10.3443 | -91.9961
10.3446 | -91.9967
10.345 | -91.9972
10.3453 | -91.9973
10.3457 | -91.9975
10.3461 | -91.9977
10.3464 | -91.997
10.3468 | -91.9971
10.3471 | -91.9975
10.3475 | -91.9981
10.3479 | -91.9983
10.3482 | -91.9987
10.3486 | -91.9996
10.349 | -92.0003
10.3493 | -92.001
10.3497 | -92.0018
10.35 | -92.0025
10.3504 | -92.0029
10.3508 | -92.0033
10.3511 | -92.0038
10.3515 | -92.004
10.3519 | -92.004
10.3522 | -92.0046
10.3526 | -92.0049
10.3529 | -92.0055
10.3533 | -92.0061
10.3537 | -92.0069
10.354 | -92.0077
10.3544 | -92.0078
10.3547 | -92.0083
10.3551 | -92.0091
10.3555 | -92.0093
10.3558 | -92.0099
10.3562 | -92.0108
10.3566 | -92.0111
10.3569 | -92.0112
10.3573 | -92.0115
10.3576 | -92.0112
10.358 | -92.011
10.3584 | -92.0115
10.3587 | -92.0123
10.3591 | -92.0132
10.3594 | -92.0132
10.3598 | -92.0132
10.3602 | -92.0137
10.3605 | -92.0142
10.3609 | -92.0147
10.3613 | -92.0154
10.3616 | -92.016
10.362 | -92.0164
10.3623 | -92.0168
10.3627 | -92.0172
10.3631 | -92.0177
10.3634 | -92.0182
10.3638 | -92.0188
10.3642 | -92.0194
10.3645 | -92.0199
10.3649 | -92.02
10.3652 | -92.0204
10.3656 | -92.0212
10.366 | -92.0211
10.3663 | -92.0211
10.3667 | -92.0217
10.367 | -92.0228
10.3674 | -92.0238
10.3678 | -92.0241
10.3681 | -92.0239
10.3685 | -92.024
10.3689 | -92.0246
10.3692 | -92.0251
10.3696 | -92.0259
10.3699 | -92.027
10.3703 | -92.0275
10.3707 | -92.0278
10.371 | -92.0282
10.3714 | -92.0289
10.3718 | -92.0295
10.3721 | -92.0303
10.3725 | -92.0307
10.3728 | -92.0312
10.3732 | -92.0315
10.3736 | -92.0318
10.3739 | -92.0323
10.3743 | -92.033
10.3746 | -92.0343
10.375 | -92.0356
10.3754 | -92.037
10.3757 | -92.0376
10.3761 | -92.0373
10.3765 | -92.0377
10.3768 | -92.038
10.3772 | -92.0386
10.3775 | -92.0393
10.3779 | -92.0404
10.3783 | -92.0412
10.3786 | -92.042
10.379 | -92.042
10.3793 | -92.042
10.3797 | -92.0422
10.3801 | -92.0427
10.3804 | -92.0431
10.3808 | -92.0433
10.3812 | -92.0431
10.3815 | -92.044
10.3819 | -92.0449
10.3822 | -92.0453
10.3826 | -92.0458
10.383 | -92.0463
10.3833 | -92.0469
10.3837 | -92.0477
10.3841 | -92.0481
10.3844 | -92.0487
10.3848 | -92.0495
10.3851 | -92.0501
10.3855 | -92.0507
10.3859 | -92.0514
10.3862 | -92.052
10.3866 | -92.0527
10.3869 | -92.0529
10.3873 | -92.0534
10.3877 | -92.0538
10.388 | -92.0543
10.3884 | -92.0546
10.3888 | -92.0547
10.3891 | -92.0552
10.3895 | -92.0559
10.3898 | -92.0563
10.3902 | -92.0565
10.3906 | -92.0567
10.3909 | -92.0571
10.3913 | -92.0576
10.3917 | -92.0578
10.392 | -92.0581
10.3924 | -92.0587
10.3927 | -92.0596
10.3931 | -92.0603
10.3935 | -92.0607
10.3938 | -92.0611
10.3942 | -92.0613
10.3945 | -92.0614
10.3949 | -92.0614
10.3953 | -92.0617
10.3956 | -92.062
10.396 | -92.062
10.3964 | -92.0622
10.3967 | -92.0628
10.3971 | -92.0635
10.3974 | -92.0636
10.3978 | -92.0631
10.3982 | -92.0626
10.3985 | -92.063
10.3989 | -92.0637
10.3992 | -92.0639
10.3996 | -92.0645
10.4 | -92.0647
10.4003 | -92.0656
10.4007 | -92.0666
10.4011 | -92.067
10.4014 | -92.0675
10.4018 | -92.0683
10.4021 | -92.0689
10.4025 | -92.0697
10.4029 | -92.0699
10.4032 | -92.0699
10.4036 | -92.0704
10.404 | -92.0708
10.4043 | -92.0711
10.4047 | -92.0717
10.405 | -92.0723
10.4054 | -92.073
10.4058 | -92.0741
10.4061 | -92.0744
10.4065 | -92.0749
10.4068 | -92.0755
10.4072 | -92.0752
10.4076 | -92.0752
10.4079 | -92.0753
10.4083 | -92.0758
10.4087 | -92.0764
10.409 | -92.0767
10.4094 | -92.0762
10.4097 | -92.0759
10.4101 | -92.0761
10.4105 | -92.0765
10.4108 | -92.0772
10.4112 | -92.0782
10.4115 | -92.0792
10.4119 | -92.0788
10.4123 | -92.0787
10.4126 | -92.0787
10.413 | -92.0788
10.4134 | -92.0788
10.4137 | -92.0795
10.4141 | -92.0801
10.4144 | -92.0804
10.4148 | -92.0809
10.4152 | -92.0811
10.4155 | -92.0814
10.4159 | -92.0821
10.4163 | -92.0823
10.4166 | -92.0823
10.417 | -92.0828
10.4173 | -92.0833
10.4177 | -92.0838
10.4181 | -92.0845
10.4184 | -92.085
10.4188 | -92.0851
10.4191 | -92.0858
10.4195 | -92.0861
10.4199 | -92.0859
10.4202 | -92.0865
10.4206 | -92.0872
10.421 | -92.0876
10.4213 | -92.0875
10.4217 | -92.0876
10.422 | -92.0883
10.4224 | -92.0887
10.4228 | -92.089
10.4231 | -92.0901
10.4235 | -92.0907
10.4238 | -92.0909
10.4242 | -92.0916
10.4246 | -92.0918
10.4249 | -92.0919
10.4253 | -92.0924
10.4257 | -92.0926
10.426 | -92.0935
10.4264 | -92.0945
10.4267 | -92.0955
10.4271 | -92.0962
10.4275 | -92.0967
10.4278 | -92.0978
10.4282 | -92.0982
10.4286 | -92.0988
10.4289 | -92.0983
10.4293 | -92.0975
10.4296 | -92.0978
10.43 | -92.0985
10.4304 | -92.0991
10.4307 | -92.0988
10.4311 | -92.099
10.4314 | -92.0995
10.4318 | -92.1001
10.4322 | -92.1007
10.4325 | -92.1009
10.4329 | -92.1012
10.4333 | -92.1018
10.4336 | -92.1027
10.434 | -92.1032
10.4343 | -92.1034
10.4347 | -92.104
10.4351 | -92.1043
10.4354 | -92.1045
10.4358 | -92.104
10.4362 | -92.1046
10.4365 | -92.1054
10.4369 | -92.1064
10.4372 | -92.1064
10.4376 | -92.1061
10.438 | -92.1062
10.4383 | -92.1065
10.4387 | -92.107
10.439 | -92.1074
10.4394 | -92.108
10.4398 | -92.1082
10.4401 | -92.1081
10.4405 | -92.1084
10.4409 | -92.1086
10.4412 | -92.1086
10.4416 | -92.1088
10.4419 | -92.1092
10.4423 | -92.1094
10.4427 | -92.1095
10.443 | -92.1095
10.4434 | -92.1093
10.4437 | -92.1097
10.4441 | -92.1106
10.4445 | -92.1107
10.4448 | -92.1108
10.4452 | -92.1111
10.4456 | -92.111
10.4459 | -92.1111
10.4463 | -92.111
10.4466 | -92.1113
10.447 | -92.1117
10.4474 | -92.112
10.4477 | -92.1122
10.4481 | -92.112
10.4485 | -92.1123
10.4488 | -92.1128
10.4492 | -92.1129
10.4495 | -92.1133
10.4499 | -92.1138
10.4503 | -92.114
10.4506 | -92.1142
10.451 | -92.1147
10.4513 | -92.115
10.4517 | -92.1151
10.4521 | -92.1155
10.4524 | -92.1158
10.4528 | -92.1158
10.4532 | -92.1162
10.4535 | -92.1167
10.4539 | -92.1167
10.4542 | -92.1169
10.4546 | -92.1181
10.455 | -92.1191
10.4553 | -92.1191
10.4557 | -92.1178
10.456 | -92.118
10.4564 | -92.1188
10.4568 | -92.1196
10.4571 | -92.12
10.4575 | -92.1201
10.4579 | -92.1201
10.4582 | -92.1204
10.4586 | -92.1208
10.4589 | -92.1217
10.4593 | -92.1219
10.4597 | -92.1218
10.46 | -92.1221
10.4604 | -92.1222
10.4608 | -92.122
10.4611 | -92.1222
10.4615 | -92.1223
10.4618 | -92.1224
10.4622 | -92.123
10.4626 | -92.1238
10.4629 | -92.124
10.4633 | -92.1241
10.4636 | -92.1237
10.464 | -92.1236
10.4644 | -92.1239
10.4647 | -92.1241
10.4651 | -92.1236
10.4655 | -92.1235
10.4658 | -92.1238
10.4662 | -92.1249
10.4665 | -92.1258
10.4669 | -92.126
10.4673 | -92.1261
10.4676 | -92.1263
10.468 | -92.1267
10.4684 | -92.1269
10.4687 | -92.1269
10.4691 | -92.1273
10.4694 | -92.1276
10.4698 | -92.128
10.4702 | -92.1269
10.4705 | -92.1268
10.4709 | -92.1269
10.4712 | -92.1271
10.4716 | -92.1277
10.472 | -92.1283
10.4723 | -92.1286
10.4727 | -92.1288
10.4731 | -92.129
10.4734 | -92.1308
10.4738 | -92.1317
10.4741 | -92.1315
10.4745 | -92.1315
10.4749 | -92.1316
10.4752 | -92.132
10.4756 | -92.1323
10.4759 | -92.1324
10.4763 | -92.1325
10.4767 | -92.1325
10.477 | -92.1325
10.4774 | -92.1324
10.4778 | -92.1325
10.4781 | -92.1334
10.4785 | -92.1341
10.4788 | -92.1338
10.4792 | -92.1349
10.4796 | -92.1367
10.4799 | -92.1375
10.4803 | -92.1379
10.4807 | -92.1381
10.481 | -92.1379
10.4814 | -92.1381
10.4817 | -92.1381
10.4821 | -92.1382
10.4825 | -92.1376
10.4828 | -92.1372
10.4832 | -92.137
10.4835 | -92.1372
10.4839 | -92.1378
10.4843 | -92.138
10.4846 | -92.1377
10.485 | -92.1379
10.4854 | -92.1373
10.4857 | -92.137
10.4861 | -92.1373
10.4864 | -92.1373
10.4868 | -92.1372
10.4872 | -92.1374
10.4875 | -92.1376
10.4879 | -92.1378
10.4883 | -92.1386
10.4886 | -92.1393
10.489 | -92.14
10.4893 | -92.1408
10.4897 | -92.1423
10.4901 | -92.1442
10.4904 | -92.1449
10.4908 | -92.1449
10.4911 | -92.1451
10.4915 | -92.1452
10.4919 | -92.1452
10.4922 | -92.1453
10.4926 | -92.1454
10.493 | -92.1453
10.4933 | -92.1452
10.4937 | -92.1453
10.494 | -92.1455
10.4944 | -92.1454
10.4948 | -92.1454
10.4951 | -92.1454
10.4955 | -92.1454
10.4958 | -92.1455
10.4962 | -92.1457
10.4966 | -92.1459
10.4969 | -92.1459
10.4973 | -92.1457
10.4977 | -92.1454
10.498 | -92.1449
10.4984 | -92.1448
10.4987 | -92.1452
10.4991 | -92.1454
10.4995 | -92.1454
10.4998 | -92.1457
10.5002 | -92.1459
10.5006 | -92.1456
10.5009 | -92.1455
10.5013 | -92.1458
10.5016 | -92.1457
10.502 | -92.1455
10.5024 | -92.1453
10.5027 | -92.1451
10.5031 | -92.1448
10.5034 | -92.1449
10.5038 | -92.1449
10.5042 | -92.1447
10.5045 | -92.1448
10.5049 | -92.1451
10.5053 | -92.1458
10.5056 | -92.147
10.506 | -92.1473
10.5063 | -92.1476
10.5067 | -92.1478
10.5071 | -92.1478
10.5074 | -92.1478
10.5078 | -92.148
10.5082 | -92.1484
10.5085 | -92.1485
10.5089 | -92.1484
10.5092 | -92.1484
10.5096 | -92.1485
10.51 | -92.1489
10.5103 | -92.1491
10.5107 | -92.149
10.511 | -92.1491
10.5114 | -92.1489
10.5118 | -92.1489
10.5121 | -92.1487
10.5125 | -92.1486
10.5129 | -92.1485
10.5132 | -92.1485
10.5136 | -92.1485
10.5139 | -92.1485
10.5143 | -92.1489
10.5147 | -92.1493
10.515 | -92.1497
10.5154 | -92.15
10.5157 | -92.1501
10.5161 | -92.15
10.5165 | -92.1504
10.5168 | -92.1511
10.5172 | -92.1514
10.5176 | -92.1512
10.5179 | -92.151
10.5183 | -92.1509
10.5186 | -92.1511
10.519 | -92.1512
10.5194 | -92.1512
10.5197 | -92.1511
10.5201 | -92.1504
10.5205 | -92.1503
10.5208 | -92.1506
10.5212 | -92.1505
10.5215 | -92.1504
10.5219 | -92.1504
10.5223 | -92.1507
10.5226 | -92.1504
10.523 | -92.1504
10.5233 | -92.1508
10.5237 | -92.1513
10.5241 | -92.152
10.5244 | -92.1528
10.5248 | -92.1533
10.5252 | -92.1536
10.5255 | -92.154
10.5259 | -92.1539
10.5262 | -92.1539
10.5266 | -92.154
10.527 | -92.1538
10.5273 | -92.1539
10.5277 | -92.154
10.528 | -92.1541
10.5284 | -92.1542
10.5288 | -92.1543
10.5291 | -92.1544
10.5295 | -92.1545
10.5299 | -92.1545
10.5302 | -92.1544
10.5306 | -92.1544
10.5309 | -92.1546
10.5313 | -92.1547
10.5317 | -92.1548
10.532 | -92.1548
10.5324 | -92.1548
10.5328 | -92.1541
10.5331 | -92.1541
10.5335 | -92.1547
10.5338 | -92.1551
10.5342 | -92.1552
10.5346 | -92.155
10.5349 | -92.155
10.5353 | -92.1555
10.5356 | -92.1562
10.536 | -92.1563
10.5364 | -92.1561
10.5367 | -92.1564
10.5371 | -92.1563
10.5375 | -92.1561
10.5378 | -92.1561
10.5382 | -92.1558
10.5385 | -92.1557
10.5389 | -92.1554
10.5393 | -92.1553
10.5396 | -92.1554
10.54 | -92.1553
10.5404 | -92.1552
10.5407 | -92.1552
10.5411 | -92.1552
10.5414 | -92.1551
10.5418 | -92.1549
10.5422 | -92.1546
10.5425 | -92.1548
10.5429 | -92.1547
10.5432 | -92.1545
10.5436 | -92.1536
10.544 | -92.1531
10.5443 | -92.1523
10.5447 | -92.1528
10.5451 | -92.1541
10.5454 | -92.1545
10.5458 | -92.1548
10.5461 | -92.155
10.5465 | -92.1553
10.5469 | -92.1552
10.5472 | -92.1551
10.5476 | -92.1544
10.5479 | -92.1538
10.5483 | -92.1534
10.5487 | -92.1535
10.549 | -92.1533
10.5494 | -92.1533
10.5498 | -92.1533
10.5501 | -92.1531
10.5505 | -92.1524
10.5508 | -92.1516
10.5512 | -92.1514
10.5516 | -92.1514
10.5519 | -92.1515
10.5523 | -92.1514
10.5527 | -92.1513
10.553 | -92.1516
10.5534 | -92.1514
10.5537 | -92.1517
10.5541 | -92.1519
10.5545 | -92.1521
10.5548 | -92.1523
10.5552 | -92.1524
10.5555 | -92.1528
10.5559 | -92.1535
10.5563 | -92.154
10.5566 | -92.154
10.557 | -92.1539
10.5574 | -92.154
10.5577 | -92.1538
10.5581 | -92.1531
10.5584 | -92.1534
10.5588 | -92.1546
10.5592 | -92.1545
10.5595 | -92.1548
10.5599 | -92.1547
10.5602 | -92.1544
10.5606 | -92.1542
10.561 | -92.1542
10.5613 | -92.1541
10.5617 | -92.154
10.5621 | -92.154
10.5624 | -92.1542
10.5628 | -92.1543
10.5631 | -92.1543
10.5635 | -92.1544
10.5639 | -92.1548
10.5642 | -92.1545
10.5646 | -92.1543
10.565 | -92.1538
10.5653 | -92.1533
10.5657 | -92.1531
10.566 | -92.1529
10.5664 | -92.1522
10.5668 | -92.1523
10.5671 | -92.1519
10.5675 | -92.1518
10.5678 | -92.1518
10.5682 | -92.1519
10.5686 | -92.1517
10.5689 | -92.1519
10.5693 | -92.1524
10.5697 | -92.1529
10.57 | -92.1529
10.5704 | -92.1526
10.5707 | -92.1523
10.5711 | -92.152
10.5715 | -92.1515
10.5718 | -92.1514
10.5722 | -92.1515
10.5725 | -92.1515
10.5729 | -92.1511
10.5733 | -92.1509
10.5736 | -92.1512
10.574 | -92.152
10.5744 | -92.1513
10.5747 | -92.1511
10.5751 | -92.151
10.5754 | -92.151
10.5758 | -92.1512
10.5762 | -92.1513
10.5765 | -92.1512
10.5769 | -92.1512
10.5773 | -92.151
10.5776 | -92.151
10.578 | -92.151
10.5783 | -92.1506
10.5787 | -92.1504
10.5791 | -92.1503
10.5794 | -92.1504
10.5798 | -92.1508
10.5801 | -92.1508
10.5805 | -92.1503
10.5809 | -92.1498
10.5812 | -92.1498
10.5816 | -92.1502
10.582 | -92.1504
10.5823 | -92.1505
10.5827 | -92.1508
10.583 | -92.1509
10.5834 | -92.1512
10.5838 | -92.1514
10.5841 | -92.1516
10.5845 | -92.1522
10.5849 | -92.1526
10.5852 | -92.1526
10.5856 | -92.1524
10.5859 | -92.153
10.5863 | -92.1532
10.5867 | -92.1535
10.587 | -92.1532
10.5874 | -92.1528
10.5877 | -92.1527
10.5881 | -92.1523
10.5885 | -92.1527
10.5888 | -92.154
10.5892 | -92.1543
10.5896 | -92.1541
10.5899 | -92.1541
10.5903 | -92.1538
10.5906 | -92.1527
10.591 | -92.1525
10.5914 | -92.1527
10.5917 | -92.1526
10.5921 | -92.1525
10.5924 | -92.1525
10.5928 | -92.1536
10.5932 | -92.1554
10.5935 | -92.1557
10.5939 | -92.1555
10.5943 | -92.1555
10.5946 | -92.1547
10.595 | -92.1534
10.5953 | -92.1522
10.5957 | -92.1526
10.5961 | -92.1529
10.5964 | -92.1526
10.5968 | -92.1529
10.5972 | -92.1526
10.5975 | -92.152
10.5979 | -92.1517
10.5982 | -92.1515
10.5986 | -92.1512
10.599 | -92.1511
10.5993 | -92.1512
10.5997 | -92.151
10.6 | -92.1511
10.6004 | -92.1519
10.6008 | -92.1512
10.6011 | -92.1506
10.6015 | -92.1504
10.6019 | -92.15
10.6022 | -92.1498
10.6026 | -92.1502
10.6029 | -92.1499
10.6033 | -92.1495
10.6037 | -92.1493
10.604 | -92.1494
10.6044 | -92.1487
10.6047 | -92.1478
10.6051 | -92.1473
10.6055 | -92.1472
10.6058 | -92.1474
10.6062 | -92.1477
10.6066 | -92.1474
10.6069 | -92.1474
10.6073 | -92.1482
10.6076 | -92.1489
10.608 | -92.1488
10.6084 | -92.1486
10.6087 | -92.1482
10.6091 | -92.1481
10.6095 | -92.1479
10.6098 | -92.1478
10.6102 | -92.1478
10.6105 | -92.1478
10.6109 | -92.1476
10.6113 | -92.1477
10.6116 | -92.1476
10.612 | -92.1474
10.6123 | -92.1473
10.6127 | -92.1475
10.6131 | -92.1475
10.6134 | -92.1473
10.6138 | -92.1473
10.6142 | -92.1473
10.6145 | -92.1472
10.6149 | -92.147
10.6152 | -92.1466
10.6156 | -92.146
10.616 | -92.145
10.6163 | -92.1455
10.6167 | -92.1461
10.6171 | -92.146
10.6174 | -92.1456
10.6178 | -92.1452
10.6181 | -92.1448
10.6185 | -92.1448
10.6189 | -92.1448
10.6192 | -92.1447
10.6196 | -92.1443
10.6199 | -92.1439
10.6203 | -92.1435
10.6207 | -92.143
10.621 | -92.1428
10.6214 | -92.1426
10.6218 | -92.1421
10.6221 | -92.1418
10.6225 | -92.1417
10.6228 | -92.1415
10.6232 | -92.1412
10.6236 | -92.1417
10.6239 | -92.1417
10.6243 | -92.1417
10.6247 | -92.1419
10.625 | -92.1418
10.6254 | -92.142
10.6257 | -92.1423
10.6261 | -92.1429
10.6265 | -92.1432
10.6268 | -92.1433
10.6272 | -92.143
10.6275 | -92.142
10.6279 | -92.1408
10.6283 | -92.1402
10.6286 | -92.1395
10.629 | -92.1389
10.6294 | -92.1384
10.6297 | -92.1378
10.6301 | -92.1379
10.6304 | -92.1376
10.6308 | -92.1375
10.6312 | -92.1374
10.6315 | -92.1376
10.6319 | -92.1378
10.6322 | -92.1382
10.6326 | -92.1382
10.633 | -92.1382
10.6333 | -92.1377
10.6337 | -92.1373
10.6341 | -92.1371
10.6344 | -92.1369
10.6348 | -92.1367
10.6351 | -92.1361
10.6355 | -92.1357
10.6359 | -92.1358
10.6362 | -92.1359
10.6366 | -92.1361
10.637 | -92.1363
10.6373 | -92.1367
10.6377 | -92.1372
10.638 | -92.1374
10.6384 | -92.1372
10.6388 | -92.1369
10.6391 | -92.1367
10.6395 | -92.1362
10.6398 | -92.136
10.6402 | -92.1361
10.6406 | -92.1355
10.6409 | -92.1352
10.6413 | -92.1353
10.6417 | -92.1352
10.642 | -92.1353
10.6424 | -92.1351
10.6427 | -92.1353
10.6431 | -92.1354
10.6435 | -92.1351
10.6438 | -92.1348
10.6442 | -92.1344
10.6445 | -92.1341
10.6449 | -92.1339
10.6453 | -92.1338
10.6456 | -92.1335
10.646 | -92.1335
10.6464 | -92.1334
10.6467 | -92.1334
10.6471 | -92.1333
10.6474 | -92.133
10.6478 | -92.1322
10.6482 | -92.1319
10.6485 | -92.1318
10.6489 | -92.1318
10.6493 | -92.1318
10.6496 | -92.1318
10.65 | -92.1319
10.6503 | -92.1319
10.6507 | -92.1317
10.6511 | -92.1316
10.6514 | -92.1314
10.6518 | -92.1303
10.6521 | -92.1292
10.6525 | -92.1289
10.6529 | -92.1289
10.6532 | -92.1292
10.6536 | -92.1292
10.654 | -92.129
10.6543 | -92.1287
10.6547 | -92.1284
10.655 | -92.1275
10.6554 | -92.1265
10.6558 | -92.126
10.6561 | -92.1259
10.6565 | -92.1259
10.6569 | -92.1262
10.6572 | -92.126
10.6576 | -92.1258
10.6579 | -92.1261
10.6583 | -92.1256
10.6587 | -92.1245
10.659 | -92.1246
10.6594 | -92.1248
10.6597 | -92.1246
10.6601 | -92.1251
10.6605 | -92.1258
10.6608 | -92.1258
10.6612 | -92.1257
10.6616 | -92.1255
10.6619 | -92.1254
10.6623 | -92.1254
10.6626 | -92.1249
10.663 | -92.1241
10.6634 | -92.1239
10.6637 | -92.1235
10.6641 | -92.1236
10.6644 | -92.1235
10.6648 | -92.1237
10.6652 | -92.1236
10.6655 | -92.1234
10.6659 | -92.1229
10.6663 | -92.122
10.6666 | -92.1212
10.667 | -92.121
10.6673 | -92.121
10.6677 | -92.1209
10.6681 | -92.121
10.6684 | -92.1209
10.6688 | -92.1212
10.6692 | -92.1211
10.6695 | -92.121
10.6699 | -92.1208
10.6702 | -92.1207
10.6706 | -92.1208
10.671 | -92.1218
10.6713 | -92.1234
10.6717 | -92.1242
10.672 | -92.1246
10.6724 | -92.1247
10.6728 | -92.1245
10.6731 | -92.1241
10.6735 | -92.1238
10.6739 | -92.1237
10.6742 | -92.1236
10.6746 | -92.1238
10.6749 | -92.1237
10.6753 | -92.1239
10.6757 | -92.1242
10.676 | -92.1246
10.6764 | -92.1238
10.6767 | -92.1234
10.6771 | -92.1235
10.6775 | -92.1239
10.6778 | -92.1239
10.6782 | -92.1237
10.6786 | -92.1225
10.6789 | -92.1223
10.6793 | -92.1223
10.6796 | -92.122
10.68 | -92.1216
10.6804 | -92.1213
10.6807 | -92.1211
10.6811 | -92.1211
10.6815 | -92.121
10.6818 | -92.1208
10.6822 | -92.121
10.6825 | -92.1216
10.6829 | -92.1217
10.6833 | -92.122
10.6836 | -92.1225
10.684 | -92.1223
10.6843 | -92.1221
10.6847 | -92.122
10.6851 | -92.1219
10.6854 | -92.1217
10.6858 | -92.1217
10.6862 | -92.1216
10.6865 | -92.1216
10.6869 | -92.1214
10.6872 | -92.1212
10.6876 | -92.1208
10.688 | -92.1203
10.6883 | -92.1203
10.6887 | -92.1199
10.689 | -92.12
10.6894 | -92.1194
10.6898 | -92.1193
10.6901 | -92.1195
10.6905 | -92.1196
10.6909 | -92.1197
10.6912 | -92.1197
10.6916 | -92.1193
10.6919 | -92.1188
10.6923 | -92.1188
10.6927 | -92.1191
10.693 | -92.1193
10.6934 | -92.1192
10.6938 | -92.1193
10.6941 | -92.1197
10.6945 | -92.1203
10.6948 | -92.1203
10.6952 | -92.1204
10.6956 | -92.1206
10.6959 | -92.1212
10.6963 | -92.1218
10.6966 | -92.1218
10.697 | -92.1218
10.6974 | -92.1219
10.6977 | -92.1219
10.6981 | -92.1213
10.6985 | -92.1212
10.6988 | -92.1211
10.6992 | -92.1208
10.6995 | -92.1205
10.6999 | -92.1205
10.7003 | -92.1203
10.7006 | -92.1202
10.701 | -92.1198
10.7014 | -92.1196
10.7017 | -92.1198
10.7021 | -92.1198
10.7024 | -92.1196
10.7028 | -92.1197
10.7032 | -92.1199
10.7035 | -92.1199
10.7039 | -92.1197
10.7042 | -92.1196
10.7046 | -92.1195
10.705 | -92.1191
10.7053 | -92.1188
10.7057 | -92.1188
10.7061 | -92.119
10.7064 | -92.1189
10.7068 | -92.1187
10.7071 | -92.1185
10.7075 | -92.1188
10.7079 | -92.1182
10.7082 | -92.1176
10.7086 | -92.1174
10.7089 | -92.1173
10.7093 | -92.1165
10.7097 | -92.1163
10.71 | -92.1167
10.7104 | -92.1166
10.7108 | -92.117
10.7111 | -92.1174
10.7115 | -92.1175
10.7118 | -92.1174
10.7122 | -92.117
10.7126 | -92.1163
10.7129 | -92.116
10.7133 | -92.1162
10.7137 | -92.1163
10.714 | -92.1159
10.7144 | -92.1157
10.7147 | -92.1157
10.7151 | -92.1153
10.7155 | -92.115
10.7158 | -92.1151
10.7162 | -92.1151
10.7165 | -92.1152
10.7169 | -92.1151
10.7173 | -92.115
10.7176 | -92.115
10.718 | -92.1149
10.7184 | -92.1149
10.7187 | -92.1148
10.7191 | -92.1145
10.7194 | -92.1143
10.7198 | -92.1141
10.7202 | -92.1133
10.7205 | -92.1133
10.7209 | -92.1143
10.7212 | -92.1155
10.7216 | -92.1164
10.722 | -92.1163
10.7223 | -92.1163
10.7227 | -92.1159
10.7231 | -92.1155
10.7234 | -92.1157
10.7238 | -92.1156
10.7241 | -92.1159
10.7245 | -92.1166
10.7249 | -92.1163
10.7252 | -92.1156
10.7256 | -92.115
10.726 | -92.1145
10.7263 | -92.1145
10.7267 | -92.1145
10.727 | -92.1151
10.7274 | -92.1163
10.7278 | -92.1176
10.7281 | -92.1179
10.7285 | -92.1178
10.7288 | -92.1178
10.7292 | -92.1177
10.7296 | -92.1176
10.7299 | -92.1173
10.7303 | -92.1179
10.7307 | -92.1187
10.731 | -92.1193
10.7314 | -92.1195
10.7317 | -92.1192
10.7321 | -92.1191
10.7325 | -92.1191
10.7328 | -92.1188
10.7332 | -92.1183
10.7336 | -92.1184
10.7339 | -92.1183
10.7343 | -92.1183
10.7346 | -92.1181
10.735 | -92.1186
10.7354 | -92.1186
10.7357 | -92.1183
10.7361 | -92.1182
10.7364 | -92.1182
10.7368 | -92.1181
10.7372 | -92.1181
10.7375 | -92.1179
10.7379 | -92.1181
10.7383 | -92.118
10.7386 | -92.1182
10.739 | -92.1187
10.7393 | -92.1187
10.7397 | -92.1187
10.7401 | -92.1186
10.7404 | -92.118
10.7408 | -92.1175
10.7411 | -92.1176
10.7415 | -92.117
10.7419 | -92.1167
10.7422 | -92.1173
10.7426 | -92.1176
10.743 | -92.1176
10.7433 | -92.1176
10.7437 | -92.1176
10.744 | -92.118
10.7444 | -92.1184
10.7448 | -92.1184
10.7451 | -92.1181
10.7455 | -92.1177
10.7459 | -92.1177
10.7462 | -92.1176
10.7466 | -92.1176
10.7469 | -92.1175
10.7473 | -92.1179
10.7477 | -92.1181
10.748 | -92.1181
10.7484 | -92.1182
10.7487 | -92.118
10.7491 | -92.1183
10.7495 | -92.1181
10.7498 | -92.1177
10.7502 | -92.1177
10.7506 | -92.118
10.7509 | -92.1189
10.7513 | -92.1193
10.7516 | -92.1191
10.752 | -92.1191
10.7524 | -92.1195
10.7527 | -92.1196
10.7531 | -92.1195
10.7535 | -92.1198
10.7538 | -92.12
10.7542 | -92.1201
10.7545 | -92.12
10.7549 | -92.1201
10.7553 | -92.1203
10.7556 | -92.1204
10.756 | -92.1203
10.7563 | -92.1202
10.7567 | -92.12
10.7571 | -92.1201
10.7574 | -92.1202
10.7578 | -92.1202
10.7582 | -92.1203
10.7585 | -92.1206
10.7589 | -92.1209
10.7592 | -92.1209
10.7596 | -92.121
10.76 | -92.1211
10.7603 | -92.1212
10.7607 | -92.1214
10.761 | -92.1214
10.7614 | -92.1214
10.7618 | -92.1216
10.7621 | -92.122
10.7625 | -92.122
10.7629 | -92.122
10.7632 | -92.1222
10.7636 | -92.1224
10.7639 | -92.1223
10.7643 | -92.1219
10.7647 | -92.1219
10.765 | -92.122
10.7654 | -92.123
10.7658 | -92.1235
10.7661 | -92.1235
10.7665 | -92.1242
10.7668 | -92.1255
10.7672 | -92.1268
10.7676 | -92.1271
10.7679 | -92.127
10.7683 | -92.1271
10.7686 | -92.1272
10.769 | -92.1272
10.7694 | -92.1271
10.7697 | -92.1271
10.7701 | -92.1271
10.7705 | -92.1274
10.7708 | -92.1274
10.7712 | -92.1274
10.7715 | -92.1275
10.7719 | -92.1278
10.7723 | -92.1281
10.7726 | -92.1284
10.773 | -92.1288
10.7734 | -92.1285
10.7737 | -92.1285
10.7741 | -92.1291
10.7744 | -92.1309
10.7748 | -92.1318
10.7752 | -92.1321
10.7755 | -92.1324
10.7759 | -92.1327
10.7762 | -92.1326
10.7766 | -92.1329
10.777 | -92.1331
10.7773 | -92.133
10.7777 | -92.1334
10.7781 | -92.1336
10.7784 | -92.1335
10.7788 | -92.1337
10.7791 | -92.1338
10.7795 | -92.1341
10.7799 | -92.1345
10.7802 | -92.1351
10.7806 | -92.1354
10.7809 | -92.1359
10.7813 | -92.1355
10.7817 | -92.1358
10.782 | -92.136
10.7824 | -92.136
10.7828 | -92.1362
10.7831 | -92.1367
10.7835 | -92.1368
10.7838 | -92.1368
10.7842 | -92.1365
10.7846 | -92.1363
10.7849 | -92.1367
10.7853 | -92.1373
10.7857 | -92.1368
10.786 | -92.1363
10.7864 | -92.1365
10.7867 | -92.1364
10.7871 | -92.1364
10.7875 | -92.1366
10.7878 | -92.1366
10.7882 | -92.1362
10.7885 | -92.1364
10.7889 | -92.137
10.7893 | -92.1373
10.7896 | -92.1373
10.79 | -92.1376
10.7904 | -92.1382
10.7907 | -92.1388
10.7911 | -92.1392
10.7914 | -92.1394
10.7918 | -92.1394
10.7922 | -92.1387
10.7925 | -92.1376
10.7929 | -92.1378
10.7932 | -92.1381
10.7936 | -92.1384
10.794 | -92.138
10.7943 | -92.1381
10.7947 | -92.1388
10.7951 | -92.1389
10.7954 | -92.1384
10.7958 | -92.1384
10.7961 | -92.1384
10.7965 | -92.1387
10.7969 | -92.139
10.7972 | -92.1395
10.7976 | -92.1399
10.798 | -92.14
10.7983 | -92.1397
10.7987 | -92.1402
10.799 | -92.1396
10.7994 | -92.1394
10.7998 | -92.1397
10.8001 | -92.1399
10.8005 | -92.1401
10.8008 | -92.1402
10.8012 | -92.1402
10.8016 | -92.1403
10.8019 | -92.1403
10.8023 | -92.1406
10.8027 | -92.1414
10.803 | -92.1419
10.8034 | -92.1422
10.8037 | -92.1425
10.8041 | -92.1431
10.8045 | -92.1432
10.8048 | -92.143
10.8052 | -92.1424
10.8056 | -92.1425
10.8059 | -92.1432
10.8063 | -92.144
10.8066 | -92.1449
10.807 | -92.1457
10.8074 | -92.1457
10.8077 | -92.1463
10.8081 | -92.1465
10.8084 | -92.1466
10.8088 | -92.1467
10.8092 | -92.1471
10.8095 | -92.1478
10.8099 | -92.148
10.8103 | -92.148
10.8106 | -92.1481
10.811 | -92.1482
10.8113 | -92.1484
10.8117 | -92.1487
10.8121 | -92.1489
10.8124 | -92.1492
10.8128 | -92.1496
10.8131 | -92.1497
10.8135 | -92.1498
10.8139 | -92.1502
10.8142 | -92.1503
10.8146 | -92.1509
10.815 | -92.1509
10.8153 | -92.1513
10.8157 | -92.1515
10.816 | -92.1517
10.8164 | -92.1518
10.8168 | -92.1522
10.8171 | -92.1532
10.8175 | -92.1537
10.8179 | -92.1536
10.8182 | -92.1535
10.8186 | -92.1533
10.8189 | -92.1531
10.8193 | -92.1534
10.8197 | -92.1537
10.82 | -92.1541
10.8204 | -92.1542
10.8207 | -92.1543
10.8211 | -92.1545
10.8215 | -92.155
10.8218 | -92.1557
10.8222 | -92.1562
10.8226 | -92.1563
10.8229 | -92.1564
10.8233 | -92.1571
10.8236 | -92.1573
10.824 | -92.1575
10.8244 | -92.1577
10.8247 | -92.1579
10.8251 | -92.1574
10.8254 | -92.1578
10.8258 | -92.1581
10.8262 | -92.1582
10.8265 | -92.1584
10.8269 | -92.1587
10.8273 | -92.1589
10.8276 | -92.1594
10.828 | -92.16
10.8283 | -92.1604
10.8287 | -92.1607
10.8291 | -92.1608
10.8294 | -92.1603
10.8298 | -92.1603
10.8302 | -92.1602
10.8305 | -92.1605
10.8309 | -92.1608
10.8312 | -92.1615
10.8316 | -92.1618
10.832 | -92.1621
10.8323 | -92.1625
10.8327 | -92.1627
10.833 | -92.1633
10.8334 | -92.1635
10.8338 | -92.1637
10.8341 | -92.1637
10.8345 | -92.1638
10.8349 | -92.164
10.8352 | -92.1644
10.8356 | -92.1647
10.8359 | -92.1649
10.8363 | -92.1653
10.8367 | -92.1656
10.837 | -92.1657
10.8374 | -92.1657
10.8377 | -92.1658
10.8381 | -92.1659
10.8385 | -92.1659
10.8388 | -92.1662
10.8392 | -92.1668
10.8396 | -92.167
10.8399 | -92.167
10.8403 | -92.1672
10.8406 | -92.1674
10.841 | -92.1678
10.8414 | -92.1681
10.8417 | -92.1683
10.8421 | -92.1684
10.8425 | -92.1688
10.8428 | -92.1696
10.8432 | -92.1705
10.8435 | -92.1711
10.8439 | -92.1714
10.8443 | -92.1717
10.8446 | -92.1719
10.845 | -92.172
10.8453 | -92.172
10.8457 | -92.1717
10.8461 | -92.1719
10.8464 | -92.1724
10.8468 | -92.1727
10.8472 | -92.173
10.8475 | -92.1733
10.8479 | -92.1738
10.8482 | -92.174
10.8486 | -92.1744
10.849 | -92.175
10.8493 | -92.1755
10.8497 | -92.1755
10.8501 | -92.1757
10.8504 | -92.176
10.8508 | -92.1768
10.8511 | -92.1775
10.8515 | -92.1778
10.8519 | -92.1781
10.8522 | -92.1785
10.8526 | -92.1788
10.8529 | -92.1789
10.8533 | -92.1792
10.8537 | -92.1796
10.854 | -92.1798
10.8544 | -92.1795
10.8548 | -92.1797
10.8551 | -92.1798
10.8555 | -92.1806
10.8558 | -92.1811
10.8562 | -92.1819
10.8566 | -92.1821
10.8569 | -92.1824
10.8573 | -92.1824
10.8576 | -92.1825
10.858 | -92.1828
10.8584 | -92.183
10.8587 | -92.1832
10.8591 | -92.1833
10.8595 | -92.1832
10.8598 | -92.1835
10.8602 | -92.1838
10.8605 | -92.1841
10.8609 | -92.1851
10.8613 | -92.1856
10.8616 | -92.186
10.862 | -92.1864
10.8624 | -92.1867
10.8627 | -92.187
10.8631 | -92.1873
10.8634 | -92.1876
10.8638 | -92.188
10.8642 | -92.1884
10.8645 | -92.1891
10.8649 | -92.1895
10.8652 | -92.1898
10.8656 | -92.19
10.866 | -92.1903
10.8663 | -92.1912
10.8667 | -92.192
10.8671 | -92.1931
10.8674 | -92.1936
10.8678 | -92.194
10.8681 | -92.1945
10.8685 | -92.1943
10.8689 | -92.1943
10.8692 | -92.1944
10.8696 | -92.1943
10.87 | -92.1942
10.8703 | -92.1945
10.8707 | -92.195
10.871 | -92.1953
10.8714 | -92.1954
10.8718 | -92.1963
10.8721 | -92.1974
10.8725 | -92.1979
10.8728 | -92.1981
10.8732 | -92.1984
10.8736 | -92.1988
10.8739 | -92.1992
10.8743 | -92.1996
10.8747 | -92.2002
10.875 | -92.2003
10.8754 | -92.2003
10.8757 | -92.2005
10.8761 | -92.2008
10.8765 | -92.2011
10.8768 | -92.2012
10.8772 | -92.2016
10.8775 | -92.202
10.8779 | -92.2024
10.8783 | -92.2027
10.8786 | -92.2033
10.879 | -92.204
10.8794 | -92.2048
10.8797 | -92.2052
10.8801 | -92.2053
10.8804 | -92.2056
10.8808 | -92.2058
10.8812 | -92.2063
10.8815 | -92.2072
10.8819 | -92.2086
10.8823 | -92.2085
10.8826 | -92.2089
10.883 | -92.2093
10.8833 | -92.2098
10.8837 | -92.2102
10.8841 | -92.2107
10.8844 | -92.2119
10.8848 | -92.2127
10.8851 | -92.2132
10.8855 | -92.2135
10.8859 | -92.2137
10.8862 | -92.2135
10.8866 | -92.2135
10.887 | -92.2139
10.8873 | -92.2144
10.8877 | -92.215
10.888 | -92.2155
10.8884 | -92.2158
10.8888 | -92.2163
10.8891 | -92.2169
10.8895 | -92.2173
10.8898 | -92.2174
10.8902 | -92.2178
10.8906 | -92.2184
10.8909 | -92.2189
10.8913 | -92.2191
10.8917 | -92.2192
10.892 | -92.2188
10.8924 | -92.219
10.8927 | -92.2189
10.8931 | -92.2192
10.8935 | -92.2197
10.8938 | -92.22
10.8942 | -92.2199
10.8946 | -92.2199
10.8949 | -92.2205
10.8953 | -92.2208
10.8956 | -92.2215
10.896 | -92.2219
10.8964 | -92.2222
10.8967 | -92.2214
10.8971 | -92.2203
10.8974 | -92.2206
10.8978 | -92.2203
10.8982 | -92.2206
10.8985 | -92.2205
10.8989 | -92.2203
10.8993 | -92.2204
10.8996 | -92.2208
10.9 | -92.2212
10.9003 | -92.2216
10.9007 | -92.2219
10.9011 | -92.2223
10.9014 | -92.2231
10.9018 | -92.2244
10.9022 | -92.2254
10.9025 | -92.2259
10.9029 | -92.2263
10.9032 | -92.2265
10.9036 | -92.2265
10.904 | -92.227
10.9043 | -92.2272
10.9047 | -92.2264
10.905 | -92.2265
10.9054 | -92.2269
10.9058 | -92.2272
10.9061 | -92.2277
10.9065 | -92.2282
10.9069 | -92.2289
10.9072 | -92.2295
10.9076 | -92.2298
10.9079 | -92.2301
10.9083 | -92.2305
10.9087 | -92.2309
10.909 | -92.2313
10.9094 | -92.2317
10.9097 | -92.2322
10.9101 | -92.2329
10.9105 | -92.2336
10.9108 | -92.2342
10.9112 | -92.2347
10.9116 | -92.2351
10.9119 | -92.2351
10.9123 | -92.2353
10.9126 | -92.2357
10.913 | -92.236
10.9134 | -92.2362
10.9137 | -92.2365
10.9141 | -92.2371
10.9145 | -92.2374
10.9148 | -92.2372
10.9152 | -92.2376
10.9155 | -92.2381
10.9159 | -92.2387
10.9163 | -92.2396
10.9166 | -92.2407
10.917 | -92.2416
10.9173 | -92.2419
10.9177 | -92.2426
10.9181 | -92.2434
10.9184 | -92.2437
10.9188 | -92.2437
10.9192 | -92.2446
10.9195 | -92.2453
10.9199 | -92.2458
10.9202 | -92.2468
10.9206 | -92.2476
10.921 | -92.2478
10.9213 | -92.2478
10.9217 | -92.248
10.9221 | -92.2483
10.9224 | -92.2487
10.9228 | -92.2494
10.9231 | -92.2502
10.9235 | -92.2506
10.9239 | -92.2511
10.9242 | -92.2516
10.9246 | -92.2519
10.9249 | -92.2516
10.9253 | -92.2519
10.9257 | -92.2525
10.926 | -92.2523
10.9264 | -92.2518
10.9268 | -92.2519
10.9271 | -92.2524
10.9275 | -92.2525
10.9278 | -92.2523
10.9282 | -92.2523
10.9286 | -92.2523
10.9289 | -92.2511
10.9293 | -92.2487
10.9296 | -92.2485
10.93 | -92.2486
10.9304 | -92.2488
10.9307 | -92.2493
10.9311 | -92.2503
10.9315 | -92.2513
10.9318 | -92.2526
10.9322 | -92.2538
10.9325 | -92.2547
10.9329 | -92.2545
10.9333 | -92.2537
10.9336 | -92.2541
10.934 | -92.2546
10.9344 | -92.2544
10.9347 | -92.2545
10.9351 | -92.2546
10.9354 | -92.2551
10.9358 | -92.2555
10.9362 | -92.2561
10.9365 | -92.2563
10.9369 | -92.2568
10.9372 | -92.2574
10.9376 | -92.2581
10.938 | -92.2585
10.9383 | -92.2589
10.9387 | -92.2594
10.9391 | -92.2597
10.9394 | -92.2601
10.9398 | -92.2606
10.9401 | -92.2615
10.9405 | -92.2625
10.9409 | -92.263
10.9412 | -92.2635
10.9416 | -92.2638
10.9419 | -92.2638
10.9423 | -92.2641
10.9427 | -92.2646
10.943 | -92.2648
10.9434 | -92.2651
10.9438 | -92.2654
10.9441 | -92.2662
10.9445 | -92.267
10.9448 | -92.267
10.9452 | -92.2667
10.9456 | -92.2672
10.9459 | -92.2675
10.9463 | -92.2673
10.9467 | -92.2678
10.947 | -92.2682
10.9474 | -92.2688
10.9477 | -92.2694
10.9481 | -92.2697
10.9485 | -92.2705
10.9488 | -92.271
10.9492 | -92.2713
10.9495 | -92.2717
10.9499 | -92.272
10.9503 | -92.2725
10.9506 | -92.2732
10.951 | -92.2736
10.9514 | -92.2746
10.9517 | -92.2756
10.9521 | -92.2767
10.9524 | -92.2783
10.9528 | -92.279
10.9532 | -92.2792
10.9535 | -92.2798
10.9539 | -92.28
10.9543 | -92.2805
10.9546 | -92.2806
10.955 | -92.2802
10.9553 | -92.2795
10.9557 | -92.2797
10.9561 | -92.2809
10.9564 | -92.2813
10.9568 | -92.2814
10.9571 | -92.2811
10.9575 | -92.2812
10.9579 | -92.2815
10.9582 | -92.2821
10.9586 | -92.2824
10.959 | -92.2821
10.9593 | -92.2822
10.9597 | -92.2827
10.96 | -92.2832
10.9604 | -92.2839
10.9608 | -92.2844
10.9611 | -92.2851
10.9615 | -92.2852
10.9618 | -92.285
10.9622 | -92.2851
10.9626 | -92.2858
10.9629 | -92.2869
10.9633 | -92.288
10.9637 | -92.2892
10.964 | -92.2897
10.9644 | -92.2901
10.9647 | -92.2905
10.9651 | -92.2909
10.9655 | -92.2912
10.9658 | -92.2916
10.9662 | -92.2917
10.9666 | -92.2918
10.9669 | -92.2921
10.9673 | -92.2925
10.9676 | -92.293
10.968 | -92.2934
10.9684 | -92.2935
10.9687 | -92.2938
10.9691 | -92.2942
10.9694 | -92.2948
10.9698 | -92.2952
10.9702 | -92.2957
10.9705 | -92.2962
10.9709 | -92.297
10.9713 | -92.2974
10.9716 | -92.2978
10.972 | -92.2986
10.9723 | -92.2994
10.9727 | -92.3003
10.9731 | -92.3013
10.9734 | -92.3022
10.9738 | -92.3026
10.9741 | -92.3032
10.9745 | -92.304
10.9749 | -92.3046
10.9752 | -92.3051
10.9756 | -92.3054
10.976 | -92.3055
10.9763 | -92.3053
10.9767 | -92.3057
10.977 | -92.3062
10.9774 | -92.3064
10.9778 | -92.3066
10.9781 | -92.3068
10.9785 | -92.307
10.9789 | -92.307
10.9792 | -92.3074
10.9796 | -92.3081
10.9799 | -92.3084
10.9803 | -92.309
10.9807 | -92.3098
10.981 | -92.31
10.9814 | -92.3099
10.9817 | -92.3102
10.9821 | -92.3107
10.9825 | -92.3112
10.9828 | -92.311
10.9832 | -92.311
10.9836 | -92.3114
10.9839 | -92.3115
10.9843 | -92.3119
10.9846 | -92.3122
10.985 | -92.3129
10.9854 | -92.3135
10.9857 | -92.3136
10.9861 | -92.3142
10.9865 | -92.3153
10.9868 | -92.3154
10.9872 | -92.3155
10.9875 | -92.316
10.9879 | -92.3164
10.9883 | -92.3168
10.9886 | -92.3171
10.989 | -92.3171
10.9893 | -92.3174
10.9897 | -92.3178
10.9901 | -92.3181
10.9904 | -92.3186
10.9908 | -92.3195
10.9912 | -92.3201
10.9915 | -92.3206
10.9919 | -92.321
10.9922 | -92.3213
10.9926 | -92.322
10.993 | -92.3236