## CALIBRATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 72 |  | CTBTO
Transducer Block | 02-90C/P | 70 |  | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
Hood | 02-92B/P | 59
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE REFERENCE
Atmospheric pressure | 1010.3 | hPa |  | PR 000 150_-_CAL
Humidity | 20.7 | %/Hr |  | Measurement device
Internal Temperature of Calibration Bench | 23.34 | °C |  | IS CALIBRATOR
AUX Sensor pressure | 1020 | hPa |  | Sensor SoH at test start
AUX Sensor Temperature | 23.9 | °C |  | OK
SENSITIVITY AT 1 Hz
Output Parameters | Theoretical Value | Measured Value @1Hz | Margin (±) | Unit | Result
Sensitivity of Pressure Derivative output | 2 | 1.99 | 0.04 | mV/Pa/sec | OK | Mode d'emploi pour créer un constat de vérification
Sensitivity of Pressure output | 20 | 20 | 0.4 | mV/Pa | OK | Cliquer sur "DEMARRER LA SAISIE"
Sensitivity of Calibration coil | 6 | 5.99 | 0.13 | Pa/V | OK | Suivre les instructions affichées
Une fois la partie SENSOR IDENTIFICATION remplie, renseigner la partie ENVIRONNEMENTAL PARAMETERS (les données environnementales)
TRANSFER FUNCTION | [5 empty] | Renseigner la partie SENSITIVITY AT 1Hz :les sensibilités en voie de dérivée de Pression (Pressure Derivative), en voie de pression (Pressure) et de Calibration (calibration Coil)
Electrical |  | Acoustical |  |  |  | Une sauvegarde temporaire est toujours possible si le fichier doit être modifié (Fichier "enregistrer")
Vérifier toutes les données saisie notamment la partie CUSTOMER ("Client") et la partie VISA (Noms des personnes et date d'ajustage)
Cliquer sur "ENREGISTREMENT DEFINITIF" (Attention le fichier sera enregistré et verrouillé dans le dossier courant, il faut retirer manuellement la protection)
SENSOR SELF NOISE |  |  |  | Decision : considering the specifications given by the manufacturer, this sensor is
CONFORM
VISA
|  |  |  | Adjusted by | A.HUE
|  |  |  | on | 2016-03-16
|  |  |  | Approved by | A.HUE
|  |  |  | Qualification | Optronic Engineer
Self Noise at 1 HZ | 6.9984e-05 | Pa/√Hz |  | on | 2016-03-17 | -83.1 | (dBPa/√Hz)
| < 0,13 | mPa/√Hz
OK
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which define the verification report. It can be used to demonstrate the link between the measure equipment and the standard national or international equipment, provided it meets the recommendations of the documentation booklet X 07-015
[8 empty] | Client | Adresse Client
[8 empty] | CEA | C.E.A. Dam. Bruyères Domaine du Grand Rué 91297 ARPAJON CEDEX FRANCE
[8 empty] | CTBTO | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
[8 empty] | IPGP | Université Paris Diderot Paris 7 Bat Lamarck UFR STEP 35 rue H.Brion 6ème étage 75205 Paris CEDEX 13
[8 empty] | SANDIA | C/O SANDIATIONAL LABS ALBUQUERQUE, NM 87123 UNITED STATES
[8 empty] | ISLA | Haiwaï Institute of Geophysics & Planetology 73-984 Makako Bay Drive KAILUA-KONA HAWAÏ
ENREGISTREMENT DEFINITIF
DEMARRER LA SAISIE

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.0797 | 11.9201
0.097656 | -34.0417 | 7.4905
0.12207 | -33.8605 | 5.4259
0.14648 | -33.9542 | 3.6126
0.1709 | -33.9074 | 3.3758
0.19531 | -33.911 | 2.8665
0.21973 | -33.9286 | 2.093
0.24414 | -33.9456 | 1.511
0.26855 | -33.937 | 1.7624
0.29297 | -33.9452 | 1.0341
0.31738 | -33.8733 | 0.71759
0.3418 | -33.92 | 0.37688
0.36621 | -33.9443 | 0.72423
0.39062 | -33.877 | 0.50366
0.41504 | -33.9553 | -0.097035
0.43945 | -33.9342 | 0.20662
0.46387 | -33.885 | -0.38867
0.48828 | -33.9105 | -0.49033
0.5127 | -33.9414 | -0.49893
0.53711 | -33.8896 | -0.19051
0.56152 | -33.9276 | -0.54097
0.58594 | -33.9205 | -1.1011
0.61035 | -33.9221 | -0.81251
0.63477 | -33.929 | -0.63243
0.65918 | -33.9533 | -1.1615
0.68359 | -33.8872 | -1.1823
0.70801 | -33.9337 | -1.2427
0.73242 | -33.9055 | -1.3233
0.75684 | -33.8971 | -1.5317
0.78125 | -33.9099 | -0.98961
0.80566 | -33.9702 | -1.5528
0.83008 | -33.9445 | -1.4092
0.85449 | -33.8784 | -1.7098
0.87891 | -33.9181 | -1.8053
0.90332 | -33.8266 | -1.6269
0.92773 | -33.9456 | -1.6624
0.95215 | -33.956 | -1.9688
0.97656 | -33.9259 | -1.9221
1.001 | -33.9154 | -2.4846
1.0254 | -33.9535 | -1.7457
1.0498 | -33.9047 | -1.893
1.0742 | -33.9553 | -2.779
1.0986 | -33.9274 | -2.4593
1.123 | -33.9434 | -2.5859
1.1475 | -33.9702 | -2.8436
1.1719 | -34.0097 | -2.5062
1.1963 | -33.9784 | -2.8158
1.2207 | -33.9704 | -2.4229
1.2451 | -34.0174 | -2.7557
1.2695 | -33.9727 | -3.0768
1.2939 | -33.9786 | -3.0252
1.3184 | -34.0189 | -2.896
1.3428 | -34.0149 | -2.9906
1.3672 | -33.9555 | -2.8573
1.3916 | -33.9825 | -3.4902
1.416 | -33.9725 | -3.1419
1.4404 | -33.9742 | -3.2791
1.4648 | -33.9644 | -3.3404
1.4893 | -33.973 | -3.4021
1.5137 | -34.0186 | -3.2284
1.5381 | -33.9665 | -3.4885
1.5625 | -33.9832 | -3.5325
1.5869 | -33.9679 | -3.9007
1.6113 | -33.9844 | -4.1075
1.6357 | -33.9271 | -4.1837
1.6602 | -33.9763 | -4.1696
1.6846 | -33.89 | -4.3531
1.709 | -33.9803 | -4.2878
1.7334 | -33.9295 | -4.1188
1.7578 | -34.0029 | -4.2761
1.7822 | -33.858 | -4.3838
1.8066 | -33.9159 | -4.3517
1.8311 | -33.9664 | -4.9423
1.8555 | -34.0551 | -4.6112
1.8799 | -33.9549 | -4.7437
1.9043 | -34.0404 | -5.1626
1.9287 | -33.972 | -4.6427
1.9531 | -34.0078 | -4.9447
1.9775 | -33.9785 | -4.659
2.002 | -33.9999 | -5.101
2.0264 | -33.9341 | -4.6991
2.0508 | -33.9501 | -5.1226
2.0752 | -33.9658 | -4.8779
2.0996 | -33.9978 | -4.9886
2.124 | -33.9757 | -4.8874
2.1484 | -33.9661 | -5.5038
2.1729 | -33.9356 | -5.6998
2.1973 | -33.9218 | -5.473
2.2217 | -34.026 | -5.5022
2.2461 | -33.9583 | -5.6293
2.2705 | -33.9976 | -5.5507
2.2949 | -33.988 | -5.8337
2.3193 | -33.9821 | -5.7361
2.3438 | -33.9949 | -6.0441
2.3682 | -33.953 | -5.8659
2.3926 | -33.9171 | -5.6456
2.417 | -33.9258 | -5.7591
2.4414 | -33.9329 | -5.6605
2.4658 | -34.0335 | -6.518
2.4902 | -34.0099 | -5.6518
2.5146 | -33.995 | -6.0785
2.5391 | -33.9521 | -6.0686
2.5635 | -33.9463 | -6.1652
2.5879 | -34.0358 | -6.6209
2.6123 | -33.9348 | -6.6673
2.6367 | -34.031 | -6.908
2.6611 | -34.0147 | -6.7669
2.6855 | -34.1026 | -6.4461
2.71 | -33.9633 | -6.6321
2.7344 | -33.9928 | -6.294
2.7588 | -34.0204 | -7.2394
2.7832 | -34.034 | -6.9983
2.8076 | -33.9794 | -6.744
2.832 | -33.9998 | -7.1101
2.8564 | -34.0194 | -7.1768
2.8809 | -34.0702 | -7.5611
2.9053 | -34.0024 | -7.5883
2.9297 | -33.9996 | -7.7068
2.9541 | -34.0041 | -7.5001
2.9785 | -34.013 | -7.4444
3.0029 | -34.0052 | -6.8156
3.0273 | -33.9688 | -7.7514
3.0518 | -34.0095 | -7.7365
3.0762 | -33.9903 | -8.2878
3.1006 | -34.01 | -7.6139
3.125 | -34.0225 | -8.0506
3.1494 | -34.0207 | -8.1668
3.1738 | -33.9876 | -8.0314
3.1982 | -34.0337 | -8.2774
3.2227 | -34.0085 | -7.8575
3.2471 | -33.9756 | -7.8366
3.2715 | -33.9804 | -7.9392
3.2959 | -34.0153 | -8.2052
3.3203 | -34.0771 | -7.9977
3.3447 | -34.0352 | -8.0578
3.3691 | -34.049 | -8.4324
3.3936 | -34.0933 | -8.5103
3.418 | -33.9655 | -7.9074
3.4424 | -34.0547 | -8.4464
3.4668 | -34.0962 | -8.4809
3.4912 | -34.0191 | -8.7396
3.5156 | -34.0452 | -8.5122
3.54 | -34.0209 | -8.8914
3.5645 | -34.037 | -8.7014
3.5889 | -33.9732 | -8.5955
3.6133 | -34.023 | -9.249
3.6377 | -34.0387 | -9.0216
3.6621 | -34.0401 | -9.0237
3.6865 | -33.9574 | -8.9777
3.7109 | -34.0975 | -9.216
3.7354 | -34.02 | -8.718
3.7598 | -34.0411 | -9.6737
3.7842 | -34.0104 | -9.2505
3.8086 | -34.0118 | -9.4312
3.833 | -33.9736 | -9.3004
3.8574 | -34.0191 | -9.6324
3.8818 | -34.049 | -9.7375
3.9062 | -34.0495 | -9.7593
3.9307 | -34.0409 | -9.3921
3.9551 | -34.0611 | -10.3082
3.9795 | -34.0739 | -9.8895
4.0039 | -34.0128 | -9.4525
4.0283 | -34.0661 | -9.9921
4.0527 | -33.986 | -10.0239
4.0771 | -34.1029 | -10.2458
4.1016 | -34.0281 | -10.1
4.126 | -34.0258 | -10.6693
4.1504 | -34.0795 | -10.406
4.1748 | -34.1123 | -10.4009
4.1992 | -34.1453 | -10.7052
4.2236 | -34.0313 | -10.7233
4.248 | -34.0579 | -10.3859
4.2725 | -34.1238 | -10.6705
4.2969 | -34.0708 | -10.1873
4.3213 | -34.0309 | -10.7139
4.3457 | -34.1313 | -11.0638
4.3701 | -34.0921 | -11.0344
4.3945 | -34.1129 | -10.9327
4.4189 | -34.0416 | -10.7548
4.4434 | -34.066 | -10.6944
4.4678 | -34.1154 | -11.2378
4.4922 | -34.0883 | -11.707
4.5166 | -34.0552 | -11.472
4.541 | -34.1072 | -11.7802
4.5654 | -34.107 | -11.6623
4.5898 | -34.2097 | -12.0947
4.6143 | -34.0926 | -11.5549
4.6387 | -34.0994 | -11.3175
4.6631 | -34.0932 | -11.5544
4.6875 | -34.059 | -12.148
4.7119 | -34.0157 | -11.7284
4.7363 | -34.0894 | -12.0705
4.7607 | -34.1149 | -12.5225
4.7852 | -34.1369 | -12.5625
4.8096 | -34.1524 | -12.1581
4.834 | -34.1247 | -12.4978
4.8584 | -34.1035 | -11.8749
4.8828 | -34.1101 | -11.8074
4.9072 | -34.0969 | -12.4419
4.9316 | -34.1014 | -12.7414
4.9561 | -34.1331 | -12.3664
4.9805 | -34.0633 | -12.4371
5.0049 | -34.1478 | -12.3852
5.0293 | -34.0844 | -12.2372
5.0537 | -34.0608 | -12.4058
5.0781 | -34.0455 | -12.4562
5.1025 | -34.0956 | -12.3161
5.127 | -34.115 | -12.5214
5.1514 | -34.0955 | -12.6804
5.1758 | -34.0969 | -12.8979
5.2002 | -34.0844 | -13.0726
5.2246 | -34.1334 | -12.7539
5.249 | -34.1121 | -13.1838
5.2734 | -34.1129 | -12.7625
5.2979 | -34.0918 | -13.3195
5.3223 | -34.182 | -13.4334
5.3467 | -34.1234 | -13.3184
5.3711 | -34.0994 | -13.6796
5.3955 | -34.1403 | -14.0527
5.4199 | -34.1534 | -13.8834
5.4443 | -34.163 | -13.5349
5.4688 | -34.1382 | -13.2097
5.4932 | -34.0548 | -13.9528
5.5176 | -34.1537 | -13.8383
5.542 | -34.1359 | -13.4194
5.5664 | -34.1115 | -14.0574
5.5908 | -34.1369 | -14.0756
5.6152 | -34.0873 | -14.2564
5.6396 | -34.0742 | -14.3101
5.6641 | -34.0755 | -14.1367
5.6885 | -34.1242 | -14.2129
5.7129 | -34.1259 | -14.2615
5.7373 | -34.115 | -13.9253
5.7617 | -34.0663 | -14.2545
5.7861 | -34.1387 | -14.2475
5.8105 | -34.1184 | -14.0543
5.835 | -34.0968 | -13.9487
5.8594 | -34.0522 | -14.6855
5.8838 | -34.1359 | -15.1462
5.9082 | -34.188 | -14.5718
5.9326 | -34.1606 | -13.8539
5.957 | -34.128 | -13.8986
5.9814 | -34.1364 | -14.6102
6.0059 | -34.1009 | -14.655
6.0303 | -34.1789 | -14.6684
6.0547 | -34.1132 | -14.745
6.0791 | -34.0989 | -14.5673
6.1035 | -34.1577 | -14.6951
6.1279 | -34.1627 | -15.1413
6.1523 | -34.2053 | -15.1755
6.1768 | -34.1677 | -14.9609
6.2012 | -34.1846 | -14.9158
6.2256 | -34.2084 | -15.4694
6.25 | -34.1176 | -15.0952
6.2744 | -34.1017 | -15.4658
6.2988 | -34.1804 | -15.7015
6.3232 | -34.133 | -15.573
6.3477 | -34.1421 | -15.427
6.3721 | -34.0666 | -15.6655
6.3965 | -34.1367 | -16.1987
6.4209 | -34.1074 | -16.3881
6.4453 | -34.0822 | -16.3175
6.4697 | -34.1301 | -16.1134
6.4941 | -34.2258 | -16.7883
6.5186 | -34.2693 | -16.4862
6.543 | -34.2103 | -16.026
6.5674 | -34.1449 | -16.7616
6.5918 | -34.1705 | -16.8021
6.6162 | -34.1791 | -16.2503
6.6406 | -34.118 | -16.3468
6.665 | -34.1591 | -16.1254
6.6895 | -34.1825 | -16.4351
6.7139 | -34.1668 | -16.6297
6.7383 | -34.1384 | -16.8251
6.7627 | -34.171 | -16.1
6.7871 | -34.1785 | -16.1315
6.8115 | -34.2014 | -16.337
6.8359 | -34.1841 | -16.5036
6.8604 | -34.1852 | -16.5386
6.8848 | -34.1815 | -16.3384
6.9092 | -34.1597 | -17.0189
6.9336 | -34.2091 | -16.3699
6.958 | -34.0955 | -16.9132
6.9824 | -34.1869 | -16.7265
7.0068 | -34.1652 | -16.8068
7.0312 | -34.1362 | -16.5921
7.0557 | -34.2047 | -17.3444
7.0801 | -34.2429 | -17.2502
7.1045 | -34.2126 | -17.308
7.1289 | -34.2938 | -17.2693
7.1533 | -34.1919 | -16.998
7.1777 | -34.2079 | -17.3796
7.2021 | -34.1793 | -17.092
7.2266 | -34.2003 | -17.5883
7.251 | -34.2229 | -18.0525
7.2754 | -34.2361 | -18.1408
7.2998 | -34.1897 | -17.7601
7.3242 | -34.0988 | -18.0047
7.3486 | -34.1921 | -18.5437
7.373 | -34.2379 | -19.0982
7.3975 | -34.1876 | -18.3958
7.4219 | -34.2981 | -19.154
7.4463 | -34.2953 | -18.0368
7.4707 | -34.2575 | -18.3184
7.4951 | -34.2417 | -17.9325
7.5195 | -34.1932 | -18.5486
7.5439 | -34.217 | -18.2709
7.5684 | -34.0954 | -18.3686
7.5928 | -34.2893 | -19.1095
7.6172 | -34.2237 | -18.7964
7.6416 | -34.2463 | -18.7402
7.666 | -34.163 | -18.3613
7.6904 | -34.2915 | -18.7447
7.7148 | -34.2863 | -18.5553
7.7393 | -34.2242 | -19.0654
7.7637 | -34.2753 | -19.1986
7.7881 | -34.294 | -19.449
7.8125 | -34.2527 | -19.132
7.8369 | -34.257 | -19.5679
7.8613 | -34.271 | -19.0574
7.8857 | -34.3395 | -18.6877
7.9102 | -34.2931 | -18.5951
7.9346 | -34.2521 | -19.02
7.959 | -34.2815 | -19.3719
7.9834 | -34.2788 | -19.2406
8.0078 | -34.3931 | -19.3802
8.0322 | -34.2459 | -19.1912
8.0566 | -34.2099 | -19.2141
8.0811 | -34.2078 | -19.6833
8.1055 | -34.3087 | -20.6813
8.1299 | -34.3227 | -19.7206
8.1543 | -34.2504 | -19.7999
8.1787 | -34.2695 | -19.8366
8.2031 | -34.294 | -19.4752
8.2275 | -34.2295 | -20.1893
8.252 | -34.3577 | -20.9077
8.2764 | -34.3436 | -20.1226
8.3008 | -34.246 | -20.0646
8.3252 | -34.3008 | -19.9337
8.3496 | -34.2787 | -20.4559
8.374 | -34.2716 | -20.6256
8.3984 | -34.2543 | -19.9221
8.4229 | -34.2285 | -19.9943
8.4473 | -34.2533 | -20.4932
8.4717 | -34.2665 | -20.2176
8.4961 | -34.2587 | -20.4814
8.5205 | -34.2561 | -20.0155
8.5449 | -34.3007 | -20.47
8.5693 | -34.282 | -20.1737
8.5938 | -34.2612 | -20.3813
8.6182 | -34.2575 | -20.985
8.6426 | -34.3144 | -21.2984
8.667 | -34.3691 | -21.132
8.6914 | -34.3687 | -20.7986
8.7158 | -34.3379 | -20.6364
8.7402 | -34.3267 | -20.564
8.7646 | -34.3066 | -21.3183
8.7891 | -34.322 | -21.1655
8.8135 | -34.2112 | -21.1245
8.8379 | -34.2402 | -21.1586
8.8623 | -34.3622 | -21.9629
8.8867 | -34.4409 | -20.9831
8.9111 | -34.2907 | -22.0644
8.9355 | -34.3337 | -21.8509
8.96 | -34.2744 | -21.6078
8.9844 | -34.3321 | -22.2948
9.0088 | -34.3429 | -22.942
9.0332 | -34.4014 | -22.6039
9.0576 | -34.3443 | -22.3049
9.082 | -34.3925 | -22.3671
9.1064 | -34.3713 | -22.536
9.1309 | -34.3424 | -22.6377
9.1553 | -34.2808 | -22.3837
9.1797 | -34.3308 | -22.9604
9.2041 | -34.3575 | -23.4543
9.2285 | -34.4299 | -22.9378
9.2529 | -34.3283 | -22.589
9.2773 | -34.3214 | -23.0926
9.3018 | -34.3495 | -23.0201
9.3262 | -34.4269 | -23.423
9.3506 | -34.3988 | -22.3892
9.375 | -34.3493 | -22.9135
9.3994 | -34.2783 | -23.2917
9.4238 | -34.3381 | -23.7365
9.4482 | -34.4784 | -23.4497
9.4727 | -34.315 | -22.5488
9.4971 | -34.3386 | -23.6657
9.5215 | -34.4262 | -23.9512
9.5459 | -34.4441 | -23.2108
9.5703 | -34.3704 | -23.1492
9.5947 | -34.3972 | -23.7579
9.6191 | -34.4051 | -22.9829
9.6436 | -34.2939 | -23.403
9.668 | -34.3194 | -24.5457
9.6924 | -34.4645 | -24.2211
9.7168 | -34.5005 | -23.8693
9.7412 | -34.4409 | -23.5198
9.7656 | -34.379 | -23.1595
9.79 | -34.3408 | -23.861
9.8145 | -34.4213 | -24.0325
9.8389 | -34.3339 | -23.7325
9.8633 | -34.3883 | -24.2458
9.8877 | -34.35 | -24.234
9.9121 | -34.3908 | -24.826
9.9365 | -34.3958 | -24.4816
9.9609 | -34.3983 | -24.1196
9.9854 | -34.3292 | -24.6208
10.0098 | -34.4835 | -25.5754

## ACOUS_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.1513 | 10.1887
0.097656 | -34.1851 | 5.4945
0.12207 | -33.9625 | 3.2381
0.14648 | -34.0927 | 1.6637
0.1709 | -33.9871 | 1.4771
0.19531 | -34.043 | 0.32012
0.21973 | -33.9946 | 0.49643
0.24414 | -34.0294 | -0.80289
0.26855 | -34.039 | -0.20174
0.29297 | -34.0048 | -1.399
0.31738 | -34.0583 | -0.95413
0.3418 | -33.9724 | -1.6882
0.36621 | -34.0726 | -1.3794
0.39062 | -33.9695 | -2.1899
0.41504 | -34.0998 | -1.7894
0.43945 | -33.9838 | -2.1168
0.46387 | -34.0862 | -2.2482
0.48828 | -33.9549 | -2.3008
0.5127 | -34.1126 | -2.5527
0.53711 | -34.004 | -2.3135
0.56152 | -34.0878 | -2.8956
0.58594 | -34.0076 | -2.415
0.61035 | -34.0819 | -3.2068
0.63477 | -34.0258 | -2.5742
0.65918 | -34.0692 | -3.3635
0.68359 | -34.0217 | -2.6209
0.70801 | -34.0232 | -3.4666
0.73242 | -34.0644 | -2.8129
0.75684 | -34.0288 | -3.6764
0.78125 | -34.0862 | -2.9957
0.80566 | -34.0085 | -3.7965
0.83008 | -34.1086 | -3.3131
0.85449 | -33.9871 | -3.7946
0.87891 | -34.0711 | -3.4759
0.90332 | -33.9723 | -3.5524
0.92773 | -34.1169 | -3.6009
0.95215 | -34.0133 | -3.5688
0.97656 | -34.1218 | -4.0778
1.001 | -34.0182 | -3.7065
1.0254 | -34.1184 | -4.251
1.0498 | -34.0579 | -3.72
1.0742 | -34.1114 | -4.6132
1.0986 | -34.0323 | -3.8862
1.123 | -34.0571 | -4.6063
1.1475 | -34.0647 | -3.8309
1.1719 | -34.0561 | -4.7552
1.1963 | -34.0879 | -3.8424
1.2207 | -34.0427 | -4.8336
1.2451 | -34.126 | -4.0845
1.2695 | -34.0193 | -4.7292
1.2939 | -34.1234 | -4.2353
1.3184 | -33.9972 | -4.6254
1.3428 | -34.1562 | -4.4052
1.3672 | -34.0278 | -4.6415
1.3916 | -34.1338 | -4.8031
1.416 | -33.9954 | -4.7115
1.4404 | -34.1277 | -4.9022
1.4648 | -34.0132 | -4.3883
1.4893 | -34.1258 | -4.926
1.5137 | -34.0567 | -4.441
1.5381 | -34.1263 | -5.3582
1.5625 | -34.0876 | -4.5427
1.5869 | -34.1063 | -5.4106
1.6113 | -34.1084 | -4.7764
1.6357 | -34.0531 | -5.6905
1.6602 | -34.0982 | -4.8872
1.6846 | -34.0425 | -5.4525
1.709 | -34.1251 | -4.9699
1.7334 | -34.0158 | -5.4818
1.7578 | -34.1583 | -5.2063
1.7822 | -34.0211 | -5.3821
1.8066 | -34.1547 | -5.3208
1.8311 | -34.0294 | -5.3092
1.8555 | -34.1837 | -5.5298
1.8799 | -34.0447 | -5.3313
1.9043 | -34.1472 | -5.9553
1.9287 | -34.0229 | -5.3995
1.9531 | -34.1271 | -6.01
1.9775 | -34.0653 | -5.1764
2.002 | -34.1142 | -6.1557
2.0264 | -34.0735 | -5.3409
2.0508 | -34.0902 | -6.0365
2.0752 | -34.1162 | -5.3862
2.0996 | -34.0973 | -6.2066
2.124 | -34.1473 | -5.5015
2.1484 | -34.0663 | -6.384
2.1729 | -34.1466 | -5.7858
2.1973 | -34.0608 | -6.1636
2.2217 | -34.1607 | -5.9564
2.2461 | -34.0611 | -6.1936
2.2705 | -34.2039 | -6.1502
2.2949 | -34.0669 | -6.2279
2.3193 | -34.1697 | -6.378
2.3438 | -34.0626 | -6.06
2.3682 | -34.1419 | -6.5581
2.3926 | -34.0662 | -5.8996
2.417 | -34.136 | -6.698
2.4414 | -34.1014 | -5.8733
2.4658 | -34.1574 | -6.7344
2.4902 | -34.1394 | -6.0025
2.5146 | -34.122 | -6.9781
2.5391 | -34.1296 | -6.2055
2.5635 | -34.1068 | -6.9427
2.5879 | -34.151 | -6.3322
2.6123 | -34.0775 | -6.7974
2.6367 | -34.169 | -6.405
2.6611 | -34.0792 | -6.8958
2.6855 | -34.202 | -6.6265
2.71 | -34.0513 | -6.877
2.7344 | -34.1953 | -6.8417
2.7588 | -34.0832 | -6.582
2.7832 | -34.1704 | -7.1198
2.8076 | -34.0724 | -6.5751
2.832 | -34.1834 | -7.211
2.8564 | -34.1258 | -6.7403
2.8809 | -34.145 | -7.5579
2.9053 | -34.1128 | -6.6912
2.9297 | -34.148 | -7.5443
2.9541 | -34.1585 | -6.894
2.9785 | -34.1138 | -7.6271
3.0029 | -34.1632 | -6.9584
3.0273 | -34.1084 | -7.6306
3.0518 | -34.2022 | -6.9307
3.0762 | -34.1083 | -7.5179
3.1006 | -34.2033 | -7.208
3.125 | -34.0897 | -7.4843
3.1494 | -34.2118 | -7.4157
3.1738 | -34.1053 | -7.5802
3.1982 | -34.2118 | -7.6281
3.2227 | -34.0732 | -7.43
3.2471 | -34.1998 | -7.8976
3.2715 | -34.1124 | -7.4416
3.2959 | -34.2015 | -7.7222
3.3203 | -34.1209 | -7.1829
3.3447 | -34.1648 | -8.0678
3.3691 | -34.1636 | -7.3538
3.3936 | -34.165 | -8.0525
3.418 | -34.1668 | -7.4236
3.4424 | -34.1418 | -8.262
3.4668 | -34.2033 | -7.5968
3.4912 | -34.1309 | -8.2884
3.5156 | -34.2355 | -7.7759
3.54 | -34.0932 | -8.2141
3.5645 | -34.2154 | -7.9843
3.5889 | -34.1213 | -8.1475
3.6133 | -34.2323 | -8.0581
3.6377 | -34.1093 | -8.171
3.6621 | -34.2293 | -8.3972
3.6865 | -34.1163 | -7.9008
3.7109 | -34.2376 | -8.3932
3.7354 | -34.1548 | -8.0088
3.7598 | -34.216 | -8.6485
3.7842 | -34.1594 | -8.1466
3.8086 | -34.1794 | -8.7681
3.833 | -34.1578 | -8.0707
3.8574 | -34.1714 | -8.7057
3.8818 | -34.2155 | -8.0866
3.9062 | -34.1506 | -8.8731
3.9307 | -34.2161 | -8.1037
3.9551 | -34.1546 | -8.7941
3.9795 | -34.2612 | -8.4301
4.0039 | -34.1323 | -8.8649
4.0283 | -34.2522 | -8.5702
4.0527 | -34.1157 | -8.7495
4.0771 | -34.2634 | -8.8626
4.1016 | -34.1566 | -8.6952
4.126 | -34.264 | -9.0509
4.1504 | -34.1479 | -8.759
4.1748 | -34.2469 | -9.3056
4.1992 | -34.1659 | -8.7519
4.2236 | -34.2009 | -9.4742
4.248 | -34.1782 | -8.6127
4.2725 | -34.1993 | -9.4488
4.2969 | -34.1832 | -8.5979
4.3213 | -34.2093 | -9.4358
4.3457 | -34.2116 | -8.7518
4.3701 | -34.1652 | -9.5285
4.3945 | -34.2343 | -8.9422
4.4189 | -34.1583 | -9.4376
4.4434 | -34.2792 | -9.011
4.4678 | -34.1838 | -9.4848
4.4922 | -34.2663 | -9.4679
4.5166 | -34.12 | -9.3951
4.541 | -34.2738 | -9.4652
4.5654 | -34.1643 | -9.2979
4.5898 | -34.2693 | -9.7663
4.6143 | -34.1684 | -9.2212
4.6387 | -34.2355 | -9.9938
4.6631 | -34.1909 | -9.2239
4.6875 | -34.2265 | -9.9995
4.7119 | -34.2081 | -9.2151
4.7363 | -34.2242 | -9.8145
4.7607 | -34.2252 | -9.2996
4.7852 | -34.2258 | -10.1743
4.8096 | -34.2571 | -9.5221
4.834 | -34.1791 | -10.1939
4.8584 | -34.3017 | -9.4425
4.8828 | -34.1773 | -9.9315
4.9072 | -34.3003 | -9.7258
4.9316 | -34.1722 | -10.1537
4.9561 | -34.2973 | -10.0584
4.9805 | -34.1704 | -9.9061
5.0049 | -34.2725 | -10.2936
5.0293 | -34.1466 | -9.9084
5.0537 | -34.2589 | -10.4233
5.0781 | -34.1813 | -9.7067
5.1025 | -34.2762 | -10.5203
5.127 | -34.2042 | -9.7919
5.1514 | -34.2712 | -10.6466
5.1758 | -34.2423 | -9.863
5.2002 | -34.2596 | -10.7626
5.2246 | -34.2636 | -9.8647
5.249 | -34.1991 | -10.642
5.2734 | -34.2741 | -10.0058
5.2979 | -34.2114 | -10.8233
5.3223 | -34.3236 | -10.2795
5.3467 | -34.1802 | -10.5807
5.3711 | -34.329 | -10.4032
5.3955 | -34.1733 | -10.618
5.4199 | -34.35 | -10.7906
5.4443 | -34.2019 | -10.5657
5.4688 | -34.2887 | -11.061
5.4932 | -34.1809 | -10.4665
5.5176 | -34.2999 | -11.0054
5.542 | -34.2076 | -10.5415
5.5664 | -34.2846 | -11.229
5.5908 | -34.2283 | -10.5324
5.6152 | -34.2692 | -11.3872
5.6396 | -34.2438 | -10.7188
5.6641 | -34.2401 | -11.4151
5.6885 | -34.2785 | -10.6902
5.7129 | -34.2081 | -11.4606
5.7373 | -34.2795 | -10.7461
5.7617 | -34.2293 | -11.2725
5.7861 | -34.3521 | -10.687
5.8105 | -34.1961 | -11.2977
5.835 | -34.3642 | -11.1503
5.8594 | -34.2039 | -11.1968
5.8838 | -34.3403 | -11.246
5.9082 | -34.2117 | -11.1997
5.9326 | -34.3297 | -11.5667
5.957 | -34.2457 | -10.9512
5.9814 | -34.3311 | -11.6479
6.0059 | -34.2422 | -11.1325
6.0303 | -34.3044 | -11.9059
6.0547 | -34.2417 | -11.1186
6.0791 | -34.2741 | -11.6202
6.1035 | -34.2979 | -10.9433
6.1279 | -34.2564 | -12.0081
6.1523 | -34.3186 | -11.3363
6.1768 | -34.2229 | -12.073
6.2012 | -34.321 | -11.5594
6.2256 | -34.2249 | -11.8317
6.25 | -34.3542 | -11.4984
6.2744 | -34.2405 | -11.7137
6.2988 | -34.3335 | -11.85
6.3232 | -34.2274 | -11.6243
6.3477 | -34.3466 | -11.6489
6.3721 | -34.2278 | -11.4451
6.3965 | -34.3536 | -12.0647
6.4209 | -34.2403 | -11.5723
6.4453 | -34.3328 | -12.0946
6.4697 | -34.2872 | -11.5631
6.4941 | -34.327 | -12.5221
6.5186 | -34.3273 | -11.6446
6.543 | -34.3125 | -12.6313
6.5674 | -34.3454 | -11.7467
6.5918 | -34.2874 | -12.698
6.6162 | -34.336 | -11.9522
6.6406 | -34.2306 | -12.532
6.665 | -34.3755 | -12.0086
6.6895 | -34.2465 | -12.3011
6.7139 | -34.3742 | -12.2977
6.7383 | -34.2327 | -12.3956
6.7627 | -34.3481 | -12.4031
6.7871 | -34.2265 | -12.1813
6.8115 | -34.4032 | -12.5283
6.8359 | -34.2855 | -12.485
6.8604 | -34.3679 | -12.8236
6.8848 | -34.255 | -12.1648
6.9092 | -34.362 | -12.8996
6.9336 | -34.3395 | -12.3474
6.958 | -34.3231 | -13.1175
6.9824 | -34.3303 | -12.2598
7.0068 | -34.3272 | -12.9796
7.0312 | -34.3591 | -12.226
7.0557 | -34.3205 | -12.9812
7.0801 | -34.3654 | -12.3212
7.1045 | -34.3233 | -12.8446
7.1289 | -34.4176 | -12.5888
7.1533 | -34.2819 | -13.0346
7.1777 | -34.3991 | -12.7776
7.2021 | -34.3047 | -12.9954
7.2266 | -34.4433 | -12.9814
7.251 | -34.3255 | -13.0237
7.2754 | -34.396 | -13.156
7.2998 | -34.2753 | -12.5264
7.3242 | -34.3772 | -13.3916
7.3486 | -34.3234 | -12.6031
7.373 | -34.3753 | -13.3812
7.3975 | -34.3612 | -12.7045
7.4219 | -34.3739 | -13.6637
7.4463 | -34.3547 | -12.7731
7.4707 | -34.3349 | -13.4741
7.4951 | -34.3588 | -13.0915
7.5195 | -34.3041 | -13.4953
7.5439 | -34.3985 | -12.997
7.5684 | -34.307 | -13.4531
7.5928 | -34.3812 | -13.1642
7.6172 | -34.289 | -13.4081
7.6416 | -34.4446 | -13.4576
7.666 | -34.2894 | -13.1962
7.6904 | -34.4535 | -13.7115
7.7148 | -34.3359 | -13.3372
7.7393 | -34.4285 | -13.8143
7.7637 | -34.3799 | -13.3372
7.7881 | -34.4255 | -14.1264
7.8125 | -34.3558 | -13.396
7.8369 | -34.3744 | -14.1006
7.8613 | -34.3831 | -13.2828
7.8857 | -34.3543 | -14.3224
7.9102 | -34.4298 | -13.3629
7.9346 | -34.3229 | -14.2865
7.959 | -34.4257 | -13.4913
7.9834 | -34.3459 | -14.1616
8.0078 | -34.4507 | -13.6337
8.0322 | -34.3602 | -14.2245
8.0566 | -34.454 | -13.9899
8.0811 | -34.3449 | -13.8064
8.1055 | -34.4631 | -13.8705
8.1299 | -34.3152 | -13.8106
8.1543 | -34.4499 | -14.3226
8.1787 | -34.3507 | -13.8292
8.2031 | -34.399 | -14.2397
8.2275 | -34.3758 | -13.8267
8.252 | -34.4158 | -14.5531
8.2764 | -34.3998 | -13.98
8.3008 | -34.4173 | -14.5696
8.3252 | -34.4201 | -13.9739
8.3496 | -34.4033 | -14.7893
8.374 | -34.4538 | -13.907
8.3984 | -34.3761 | -14.9385
8.4229 | -34.4385 | -14.1175
8.4473 | -34.3788 | -14.7007
8.4717 | -34.4479 | -14.2129
8.4961 | -34.3489 | -14.7637
8.5205 | -34.4796 | -14.1733
8.5449 | -34.3523 | -14.644
8.5693 | -34.5024 | -14.5172
8.5938 | -34.3636 | -14.2756
8.6182 | -34.4983 | -14.9459
8.6426 | -34.3835 | -14.2774
8.667 | -34.447 | -15.0213
8.6914 | -34.402 | -14.0528
8.7158 | -34.4182 | -15.2398
8.7402 | -34.4105 | -14.352
8.7646 | -34.4171 | -15.0963
8.7891 | -34.459 | -14.5432
8.8135 | -34.4258 | -15.3265
8.8379 | -34.5052 | -14.405
8.8623 | -34.4145 | -15.0527
8.8867 | -34.5102 | -14.8398
8.9111 | -34.3796 | -14.964
8.9355 | -34.4848 | -15.0622
8.96 | -34.4012 | -15.1902
8.9844 | -34.5165 | -15.1588
9.0088 | -34.4041 | -15.1377
9.0332 | -34.5073 | -15.1553
9.0576 | -34.3622 | -14.9273
9.082 | -34.542 | -15.4041
9.1064 | -34.439 | -14.7723
9.1309 | -34.4928 | -15.5929
9.1553 | -34.4039 | -14.8923
9.1797 | -34.4578 | -15.6689
9.2041 | -34.4504 | -14.9904
9.2285 | -34.4412 | -15.6941
9.2529 | -34.4427 | -14.655
9.2773 | -34.426 | -15.8444
9.3018 | -34.5155 | -15.048
9.3262 | -34.4594 | -15.7646
9.3506 | -34.5397 | -15.2983
9.375 | -34.4012 | -15.5702
9.3994 | -34.514 | -15.1786
9.4238 | -34.4302 | -15.4971
9.4482 | -34.5698 | -15.3252
9.4727 | -34.419 | -15.5647
9.4971 | -34.5224 | -15.8958
9.5215 | -34.4032 | -15.4181
9.5459 | -34.4977 | -16.1004
9.5703 | -34.4489 | -15.2311
9.5947 | -34.4833 | -16.2601
9.6191 | -34.4448 | -15.3872
9.6436 | -34.5265 | -16.1231
9.668 | -34.4955 | -15.3703
9.6924 | -34.4435 | -16.1711
9.7168 | -34.5142 | -15.5842
9.7412 | -34.4538 | -16.367
9.7656 | -34.5393 | -15.4636
9.79 | -34.4186 | -16.2398
9.8145 | -34.5487 | -15.6097
9.8389 | -34.4279 | -16.5554
9.8633 | -34.5556 | -15.9139
9.8877 | -34.4119 | -16.0419
9.9121 | -34.5513 | -15.9713
9.9365 | -34.4298 | -16.1723
9.9609 | -34.5697 | -16.1495
9.9854 | -34.466 | -15.9938
10.0098 | -34.5895 | -16.657

## SEALED_NOISE
dBPa/√Hz | Frequency
0.009216 | -35.8199
0.009545 | -36.9688
0.009874 | -37.1116
0.010203 | -38.8751
0.010532 | -39.2624
0.010862 | -37.8881
0.011191 | -34.7203
0.01152 | -33.06
0.011849 | -32.8061
0.012178 | -33.09
0.012507 | -34.919
0.012836 | -36.138
0.013166 | -35.32
0.013495 | -33.089
0.013824 | -33.0757
0.014153 | -33.3685
0.014482 | -34.5685
0.014811 | -37.6326
0.01514 | -37.6407
0.015469 | -34.4253
0.015799 | -33.3559
0.016128 | -35.8363
0.016457 | -39.6201
0.016786 | -41.0942
0.017115 | -40.5135
0.017444 | -40.86
0.017773 | -42.2068
0.018103 | -42.443
0.018432 | -41.363
0.018761 | -37.2793
0.01909 | -35.2339
0.019419 | -34.4218
0.019748 | -34.7102
0.020077 | -35.4489
0.020407 | -36.1101
0.020736 | -37.0538
0.021065 | -38.3555
0.021394 | -38.5898
0.021723 | -37.5967
0.022052 | -36.8208
0.022381 | -37.0241
0.022711 | -38.1541
0.02304 | -38.8168
0.023369 | -38.1707
0.023698 | -38.0651
0.024027 | -39.169
0.024356 | -39.9461
0.024685 | -39.2626
0.025014 | -38.5871
0.025344 | -38.4786
0.025673 | -38.9441
0.026002 | -39.6688
0.026331 | -39.9678
0.02666 | -40.2272
0.026989 | -41.2617
0.027318 | -42.1699
0.027648 | -40.9421
0.027977 | -38.7864
0.028306 | -37.9864
0.028635 | -37.7458
0.028964 | -38.1018
0.029293 | -38.917
0.029622 | -39.9873
0.029952 | -41.0706
0.030281 | -41.8105
0.03061 | -42.0437
0.030939 | -41.8674
0.031268 | -41.7281
0.031597 | -42.0395
0.031926 | -42.9715
0.032256 | -44.1698
0.032585 | -45.1866
0.032914 | -45.9354
0.033243 | -46.6337
0.033572 | -47.0767
0.033901 | -46.6743
0.03423 | -45.4488
0.034559 | -44.0339
0.034889 | -42.5808
0.035218 | -41.4733
0.035547 | -41.0853
0.035876 | -41.5857
0.036205 | -42.7337
0.036534 | -43.8321
0.036863 | -44.4564
0.037193 | -44.8492
0.037522 | -45.1088
0.037851 | -45.0557
0.03818 | -45.0154
0.038509 | -45.0508
0.038838 | -45.256
0.039167 | -45.5449
0.039497 | -45.8844
0.039826 | -46.3523
0.040155 | -47.0107
0.040484 | -47.8961
0.040813 | -48.8256
0.041142 | -49.5913
0.041471 | -50.1367
0.041801 | -50.3855
0.04213 | -50.3202
0.042459 | -49.9364
0.042788 | -49.445
0.043117 | -48.9401
0.043446 | -48.4433
0.043775 | -48.0327
0.044104 | -47.5118
0.044434 | -46.8854
0.044763 | -46.3998
0.045092 | -46.3457
0.045421 | -46.8924
0.04575 | -48.0477
0.046079 | -49.6487
0.046408 | -51.2733
0.046738 | -52.4104
0.047067 | -52.8227
0.047396 | -51.9899
0.047725 | -51.5132
0.048054 | -51.1181
0.048383 | -50.9999
0.048712 | -51.1011
0.049042 | -51.3727
0.049371 | -51.5241
0.0497 | -51.3911
0.050029 | -51.0068
0.050358 | -50.4877
0.050687 | -50.1078
0.051016 | -49.9367
0.051346 | -50.0483
0.051675 | -50.3523
0.052004 | -50.6542
0.052333 | -50.8251
0.052662 | -50.7364
0.052991 | -50.3164
0.05332 | -49.6667
0.053649 | -49.0915
0.053979 | -48.8143
0.054308 | -48.8877
0.054637 | -49.3429
0.054966 | -50.0529
0.055295 | -50.7942
0.055624 | -51.3146
0.055953 | -51.5145
0.056283 | -51.5174
0.056612 | -51.4298
0.056941 | -51.4016
0.05727 | -51.2125
0.057599 | -51.0135
0.057928 | -50.7296
0.058257 | -50.3659
0.058587 | -50.0036
0.058916 | -49.721
0.059245 | -49.5231
0.059574 | -49.4236
0.059903 | -49.4812
0.060232 | -49.7118
0.060561 | -50.0667
0.060891 | -50.4707
0.06122 | -50.8385
0.061549 | -51.1479
0.061878 | -51.4186
0.062207 | -51.659
0.062536 | -51.867
0.062865 | -52.0804
0.063194 | -52.2216
0.063524 | -52.2444
0.063853 | -52.1343
0.064182 | -51.8813
0.064511 | -51.5598
0.06484 | -51.2734
0.065169 | -51.1019
0.065498 | -51.1663
0.065828 | -51.4881
0.066157 | -52.2309
0.066486 | -52.9173
0.066815 | -53.7449
0.067144 | -54.6891
0.067473 | -55.7369
0.067802 | -56.6797
0.068132 | -57.5968
0.068461 | -58.2989
0.06879 | -58.5531
0.069119 | -58.4133
0.069448 | -58.2051
0.069777 | -57.9026
0.070106 | -57.4553
0.070436 | -56.9704
0.070765 | -56.5371
0.071094 | -56.1608
0.071423 | -55.8887
0.071752 | -55.7918
0.072081 | -55.8228
0.07241 | -55.9462
0.072739 | -56.0762
0.073069 | -56.1923
0.073398 | -56.2503
0.073727 | -56.161
0.074056 | -55.8837
0.074385 | -55.567
0.074714 | -55.2798
0.075043 | -54.9577
0.075373 | -54.6916
0.075702 | -54.6494
0.076031 | -54.5784
0.07636 | -54.6115
0.076689 | -54.717
0.077018 | -54.9338
0.077347 | -55.2343
0.077677 | -55.6054
0.078006 | -56.0592
0.078335 | -56.5799
0.078664 | -57.1533
0.078993 | -57.606
0.079322 | -57.9689
0.079651 | -58.2489
0.079981 | -58.4227
0.08031 | -58.314
0.080639 | -58.0375
0.080968 | -57.7121
0.081297 | -57.3564
0.081626 | -57.037
0.081955 | -56.7982
0.082284 | -56.701
0.082614 | -56.7453
0.082943 | -56.9456
0.083272 | -57.2586
0.083601 | -57.6718
0.08393 | -58.1208
0.084259 | -58.5615
0.084588 | -58.9571
0.084918 | -59.3057
0.085247 | -59.5832
0.085576 | -60.0156
0.085905 | -60.4093
0.086234 | -60.7948
0.086563 | -61.1626
0.086892 | -61.5215
0.087222 | -61.7252
0.087551 | -61.8518
0.08788 | -61.7972
0.088209 | -61.5889
0.088538 | -61.2503
0.088867 | -60.8684
0.089196 | -60.4706
0.089526 | -60.1125
0.089855 | -59.8172
0.090184 | -59.5533
0.090513 | -59.3674
0.090842 | -59.2051
0.091171 | -59.1171
0.0915 | -59.0752
0.091829 | -59.0544
0.092159 | -59.0231
0.092488 | -59.0041
0.092817 | -58.992
0.093146 | -58.9301
0.093475 | -58.8313
0.093804 | -58.6958
0.094133 | -58.5769
0.094463 | -58.4763
0.094792 | -58.5772
0.095121 | -58.6403
0.09545 | -58.7823
0.095779 | -58.9789
0.096108 | -59.2092
0.096437 | -59.4794
0.096767 | -59.7707
0.097096 | -60.0703
0.097425 | -60.3779
0.097754 | -60.6844
0.098083 | -61.0266
0.098412 | -61.3114
0.098741 | -61.6146
0.099071 | -61.9019
0.0994 | -62.1547
0.099729 | -62.3246
0.10006 | -62.4154
0.10039 | -62.4539
0.10072 | -62.42
0.10104 | -62.2612
0.10137 | -62.0409
0.1017 | -61.6957
0.10203 | -61.3067
0.10236 | -60.9387
0.10269 | -60.5709
0.10302 | -60.2391
0.10335 | -59.9808
0.10368 | -59.7824
0.10401 | -59.6377
0.10434 | -59.7202
0.10467 | -59.7337
0.105 | -59.826
0.10532 | -59.9947
0.10565 | -60.2312
0.10598 | -60.5442
0.10631 | -60.9498
0.10664 | -61.4204
0.10697 | -61.9467
0.1073 | -62.4467
0.10763 | -62.9339
0.10796 | -63.3279
0.10829 | -63.5871
0.10862 | -63.686
0.10895 | -63.6009
0.10927 | -63.3666
0.1096 | -63.0906
0.10993 | -62.741
0.11026 | -62.3924
0.11059 | -62.1325
0.11092 | -61.9464
0.11125 | -61.844
0.11158 | -61.8514
0.11191 | -61.9739
0.11224 | -62.1852
0.11256 | -62.4844
0.11289 | -62.8603
0.11322 | -63.2388
0.11355 | -63.627
0.11388 | -63.853
0.11421 | -64.1563
0.11454 | -64.4185
0.11487 | -64.5998
0.1152 | -64.6987
0.11553 | -64.7804
0.11586 | -64.8361
0.11619 | -64.7476
0.11651 | -64.6476
0.11684 | -64.5344
0.11717 | -64.4198
0.1175 | -64.2881
0.11783 | -64.1417
0.11816 | -63.9643
0.11849 | -63.8019
0.11882 | -63.6468
0.11915 | -63.4804
0.11948 | -63.3035
0.11981 | -63.154
0.12014 | -63.0377
0.12046 | -62.94
0.12079 | -62.8743
0.12112 | -62.8456
0.12145 | -62.8347
0.12178 | -62.8638
0.12211 | -62.9175
0.12244 | -62.9902
0.12277 | -63.0759
0.1231 | -63.1886
0.12343 | -63.3523
0.12376 | -63.4927
0.12409 | -63.647
0.12441 | -63.8286
0.12474 | -63.9965
0.12507 | -64.1246
0.1254 | -64.223
0.12573 | -64.2895
0.12606 | -64.3299
0.12639 | -64.3082
0.12672 | -64.3029
0.12705 | -64.2442
0.12738 | -64.1792
0.12771 | -64.0723
0.12804 | -63.9362
0.12836 | -63.8272
0.12869 | -63.7256
0.12902 | -63.6101
0.12935 | -63.4879
0.12968 | -63.3742
0.13001 | -63.2708
0.13034 | -63.1951
0.13067 | -63.1555
0.131 | -63.1381
0.13133 | -63.0995
0.13165 | -63.0775
0.13198 | -63.0881
0.13231 | -63.1097
0.13264 | -63.1861
0.13297 | -63.2317
0.1333 | -63.3175
0.13363 | -63.4138
0.13396 | -63.5389
0.13429 | -63.6987
0.13462 | -63.8988
0.13495 | -64.1593
0.13528 | -64.4513
0.13561 | -64.7646
0.13593 | -65.0915
0.13626 | -65.4741
0.13659 | -65.8439
0.13692 | -66.2497
0.13725 | -66.6196
0.13758 | -66.9938
0.13791 | -67.3788
0.13824 | -67.7042
0.13857 | -68.0212
0.1389 | -68.3145
0.13922 | -68.5556
0.13955 | -68.7624
0.13988 | -68.9571
0.14021 | -69.0534
0.14054 | -69.1019
0.14087 | -69.1517
0.1412 | -69.2346
0.14153 | -69.2543
0.14186 | -69.2281
0.14219 | -69.2374
0.14252 | -69.2201
0.14285 | -69.135
0.14317 | -69.0197
0.1435 | -68.8819
0.14383 | -68.6283
0.14416 | -68.3582
0.14449 | -68.0704
0.14482 | -67.7568
0.14515 | -67.41
0.14548 | -67.0432
0.14581 | -66.6881
0.14614 | -66.3458
0.14647 | -66.0279
0.1468 | -65.7493
0.14713 | -65.5118
0.14745 | -65.3257
0.14778 | -65.1862
0.14811 | -65.0905
0.14844 | -65.0494
0.14877 | -65.074
0.1491 | -65.1558
0.14943 | -65.2938
0.14976 | -65.4838
0.15009 | -65.7245
0.15042 | -66.0266
0.15074 | -66.3601
0.15107 | -66.7512
0.1514 | -67.1657
0.15173 | -67.4648
0.15206 | -67.8507
0.15239 | -68.2232
0.15272 | -68.581
0.15305 | -68.8688
0.15338 | -69.0884
0.15371 | -69.2569
0.15404 | -69.3876
0.15437 | -69.4441
0.15469 | -69.4338
0.15502 | -69.4225
0.15535 | -69.354
0.15568 | -69.219
0.15601 | -69.0913
0.15634 | -69.0005
0.15667 | -68.9151
0.157 | -68.8551
0.15733 | -68.8229
0.15766 | -68.8035
0.15799 | -68.8081
0.15832 | -68.7968
0.15864 | -68.8111
0.15897 | -68.8518
0.1593 | -68.8585
0.15963 | -68.8506
0.15996 | -68.8565
0.16029 | -68.8519
0.16062 | -68.8245
0.16095 | -68.7893
0.16128 | -68.6951
0.16161 | -68.6515
0.16194 | -68.5712
0.16226 | -68.5002
0.16259 | -68.394
0.16292 | -68.2959
0.16325 | -68.215
0.16358 | -68.1295
0.16391 | -68.0432
0.16424 | -67.9806
0.16457 | -67.9306
0.1649 | -67.8772
0.16523 | -67.8605
0.16556 | -67.8754
0.16589 | -67.9113
0.16622 | -67.9586
0.16654 | -68.0089
0.16687 | -68.0859
0.1672 | -68.169
0.16753 | -68.2806
0.16786 | -68.3982
0.16819 | -68.5227
0.16852 | -68.6639
0.16885 | -68.7864
0.16918 | -68.9143
0.16951 | -69.0608
0.16984 | -69.1834
0.17016 | -69.3398
0.17049 | -69.4688
0.17082 | -69.579
0.17115 | -69.6905
0.17148 | -69.8355
0.17181 | -69.9636
0.17214 | -70.0683
0.17247 | -70.1585
0.1728 | -70.2641
0.17313 | -70.3701
0.17346 | -70.4583
0.17378 | -70.5384
0.17411 | -70.6275
0.17444 | -70.7149
0.17477 | -70.7993
0.1751 | -70.8888
0.17543 | -70.9937
0.17576 | -71.0713
0.17609 | -71.1363
0.17642 | -71.2177
0.17675 | -71.2988
0.17708 | -71.3638
0.17741 | -71.4498
0.17773 | -71.4912
0.17806 | -71.5162
0.17839 | -71.5512
0.17872 | -71.5811
0.17905 | -71.6057
0.17938 | -71.6255
0.17971 | -71.6267
0.18004 | -71.6048
0.18037 | -71.5063
0.1807 | -71.4515
0.18103 | -71.4059
0.18135 | -71.3566
0.18168 | -71.3096
0.18201 | -71.2477
0.18234 | -71.1697
0.18267 | -71.1124
0.183 | -71.0478
0.18333 | -70.966
0.18366 | -70.912
0.18399 | -70.8835
0.18432 | -70.8622
0.18465 | -70.8454
0.18498 | -70.8581
0.1853 | -70.8668
0.18563 | -70.899
0.18596 | -70.9204
0.18629 | -70.9544
0.18662 | -70.9883
0.18695 | -71.0286
0.18728 | -71.0723
0.18761 | -71.1337
0.18794 | -71.1811
0.18827 | -71.2164
0.1886 | -71.2592
0.18893 | -71.2897
0.18925 | -71.3112
0.18958 | -71.3549
0.18991 | -71.3915
0.19024 | -71.4359
0.19057 | -71.4908
0.1909 | -71.5311
0.19123 | -71.5305
0.19156 | -71.5275
0.19189 | -71.5319
0.19222 | -71.5328
0.19255 | -71.5312
0.19287 | -71.5043
0.1932 | -71.4625
0.19353 | -71.4194
0.19386 | -71.3684
0.19419 | -71.3012
0.19452 | -71.2185
0.19485 | -71.1228
0.19518 | -71.0309
0.19551 | -70.9437
0.19584 | -70.8292
0.19617 | -70.7482
0.1965 | -70.675
0.19682 | -70.6189
0.19715 | -70.6007
0.19748 | -70.5565
0.19781 | -70.5071
0.19814 | -70.4963
0.19847 | -70.5078
0.1988 | -70.5195
0.19913 | -70.5909
0.19946 | -70.6211
0.19979 | -70.6548
0.20012 | -70.7071
0.20045 | -70.7629
0.20077 | -70.833
0.2011 | -70.9059
0.20143 | -70.9756
0.20176 | -71.049
0.20209 | -71.1355
0.20242 | -71.2298
0.20275 | -71.3136
0.20308 | -71.4164
0.20341 | -71.5195
0.20374 | -71.5754
0.20407 | -71.6329
0.20439 | -71.7097
0.20472 | -71.7719
0.20505 | -71.8401
0.20538 | -71.9047
0.20571 | -71.9581
0.20604 | -71.9915
0.20637 | -71.9931
0.2067 | -71.9576
0.20703 | -71.8995
0.20736 | -71.8371
0.20769 | -71.7636
0.20802 | -71.6816
0.20834 | -71.5916
0.20867 | -71.4622
0.209 | -71.3683
0.20933 | -71.2674
0.20966 | -71.1657
0.20999 | -71.0701
0.21032 | -70.9929
0.21065 | -70.9489
0.21098 | -70.8966
0.21131 | -70.8561
0.21164 | -70.8146
0.21196 | -70.7872
0.21229 | -70.7949
0.21262 | -70.8233
0.21295 | -70.8522
0.21328 | -70.8952
0.21361 | -70.9563
0.21394 | -71.0122
0.21427 | -71.0765
0.2146 | -71.1521
0.21493 | -71.252
0.21526 | -71.3473
0.21558 | -71.4447
0.21591 | -71.5262
0.21624 | -71.5994
0.21657 | -71.6893
0.2169 | -71.7812
0.21723 | -71.8601
0.21756 | -71.9233
0.21789 | -71.9926
0.21822 | -72.0166
0.21855 | -72.0606
0.21888 | -72.1045
0.21921 | -72.1601
0.21954 | -72.2068
0.21986 | -72.1909
0.22019 | -72.161
0.22052 | -72.1425
0.22085 | -72.1354
0.22118 | -72.1302
0.22151 | -72.1136
0.22184 | -72.077
0.22217 | -72.0484
0.2225 | -72.0226
0.22283 | -71.9927
0.22316 | -71.9726
0.22348 | -71.9523
0.22381 | -71.9419
0.22414 | -71.9456
0.22447 | -71.9595
0.2248 | -71.9785
0.22513 | -71.9951
0.22546 | -72.0183
0.22579 | -72.0513
0.22612 | -72.0819
0.22645 | -72.1037
0.22678 | -72.1403
0.22711 | -72.1901
0.22743 | -72.2695
0.22776 | -72.367
0.22809 | -72.4249
0.22842 | -72.4511
0.22875 | -72.4936
0.22908 | -72.5432
0.22941 | -72.5994
0.22974 | -72.642
0.23007 | -72.6783
0.2304 | -72.7156
0.23073 | -72.7255
0.23106 | -72.7088
0.23138 | -72.693
0.23171 | -72.6702
0.23204 | -72.6443
0.23237 | -72.6008
0.2327 | -72.5404
0.23303 | -72.4687
0.23336 | -72.3916
0.23369 | -72.3234
0.23402 | -72.2429
0.23435 | -72.1573
0.23467 | -72.0672
0.235 | -71.989
0.23533 | -71.922
0.23566 | -71.8634
0.23599 | -71.8005
0.23632 | -71.7742
0.23665 | -71.7505
0.23698 | -71.715
0.23731 | -71.7509
0.23764 | -71.7377
0.23797 | -71.7466
0.2383 | -71.7511
0.23863 | -71.7574
0.23895 | -71.7805
0.23928 | -71.8198
0.23961 | -71.8602
0.23994 | -71.8994
0.24027 | -71.9432
0.2406 | -71.9799
0.24093 | -72.0212
0.24126 | -72.0603
0.24159 | -72.1081
0.24192 | -72.17
0.24225 | -72.2313
0.24258 | -72.267
0.2429 | -72.2959
0.24323 | -72.3441
0.24356 | -72.372
0.24389 | -72.3896
0.24422 | -72.3945
0.24455 | -72.375
0.24488 | -72.3448
0.24521 | -72.3388
0.24554 | -72.3339
0.24587 | -72.3035
0.24619 | -72.2643
0.24652 | -72.218
0.24685 | -72.1889
0.24718 | -72.1589
0.24751 | -72.1403
0.24784 | -72.1193
0.24817 | -72.103
0.2485 | -72.0979
0.24883 | -72.0895
0.24916 | -72.0698
0.24949 | -72.0484
0.24982 | -72.0556
0.25015 | -72.0629
0.25047 | -72.065
0.2508 | -72.0971
0.25113 | -72.1232
0.25146 | -72.1599
0.25179 | -72.2009
0.25212 | -72.2509
0.25245 | -72.3068
0.25278 | -72.3593
0.25311 | -72.3989
0.25344 | -72.4412
0.25377 | -72.4883
0.25409 | -72.5329
0.25442 | -72.5768
0.25475 | -72.6187
0.25508 | -72.6577
0.25541 | -72.7023
0.25574 | -72.7431
0.25607 | -72.7872
0.2564 | -72.8031
0.25673 | -72.8583
0.25706 | -72.9014
0.25739 | -72.9386
0.25772 | -72.9768
0.25804 | -73.0161
0.25837 | -73.061
0.2587 | -73.1218
0.25903 | -73.1714
0.25936 | -73.2293
0.25969 | -73.2804
0.26002 | -73.3183
0.26035 | -73.3551
0.26068 | -73.412
0.26101 | -73.4682
0.26134 | -73.5168
0.26166 | -73.568
0.26199 | -73.6022
0.26232 | -73.6536
0.26265 | -73.7217
0.26298 | -73.7659
0.26331 | -73.7995
0.26364 | -73.8381
0.26397 | -73.8701
0.2643 | -73.8924
0.26463 | -73.8922
0.26496 | -73.8949
0.26528 | -73.9041
0.26561 | -73.9018
0.26594 | -73.9062
0.26627 | -73.918
0.2666 | -73.9154
0.26693 | -73.9125
0.26726 | -73.9103
0.26759 | -73.9139
0.26792 | -73.9422
0.26825 | -73.9931
0.26858 | -74.0143
0.26891 | -74.0254
0.26924 | -74.0329
0.26956 | -74.0359
0.26989 | -74.0547
0.27022 | -74.0771
0.27055 | -74.1178
0.27088 | -74.1324
0.27121 | -74.1507
0.27154 | -74.1896
0.27187 | -74.2394
0.2722 | -74.297
0.27253 | -74.3492
0.27286 | -74.402
0.27318 | -74.4561
0.27351 | -74.514
0.27384 | -74.5719
0.27417 | -74.6237
0.2745 | -74.6799
0.27483 | -74.7356
0.27516 | -74.7848
0.27549 | -74.8617
0.27582 | -74.9185
0.27615 | -74.963
0.27648 | -75.0004
0.27681 | -75.0526
0.27713 | -75.1034
0.27746 | -75.1593
0.27779 | -75.206
0.27812 | -75.2505
0.27845 | -75.2773
0.27878 | -75.2922
0.27911 | -75.3195
0.27944 | -75.3638
0.27977 | -75.4167
0.2801 | -75.4482
0.28043 | -75.4876
0.28075 | -75.5211
0.28108 | -75.5579
0.28141 | -75.6127
0.28174 | -75.6198
0.28207 | -75.6263
0.2824 | -75.6621
0.28273 | -75.6659
0.28306 | -75.6624
0.28339 | -75.6615
0.28372 | -75.6397
0.28405 | -75.6381
0.28437 | -75.6138
0.2847 | -75.5438
0.28503 | -75.4997
0.28536 | -75.4496
0.28569 | -75.4073
0.28602 | -75.3667
0.28635 | -75.3209
0.28668 | -75.277
0.28701 | -75.2255
0.28734 | -75.1689
0.28767 | -75.1167
0.288 | -75.058
0.28832 | -74.999
0.28865 | -74.9357
0.28898 | -74.8734
0.28931 | -74.8299
0.28964 | -74.7762
0.28997 | -74.7172
0.2903 | -74.6627
0.29063 | -74.6165
0.29096 | -74.5931
0.29129 | -74.5668
0.29162 | -74.5212
0.29195 | -74.4822
0.29227 | -74.4538
0.2926 | -74.4313
0.29293 | -74.4111
0.29326 | -74.3898
0.29359 | -74.3819
0.29392 | -74.3739
0.29425 | -74.3991
0.29458 | -74.3946
0.29491 | -74.3891
0.29524 | -74.4057
0.29557 | -74.4191
0.2959 | -74.4352
0.29622 | -74.4547
0.29655 | -74.4792
0.29688 | -74.5111
0.29721 | -74.5501
0.29754 | -74.5761
0.29787 | -74.606
0.2982 | -74.6373
0.29853 | -74.6791
0.29886 | -74.7216
0.29919 | -74.7824
0.29952 | -74.8484
0.29984 | -74.9016
0.30017 | -74.952
0.3005 | -75.0061
0.30083 | -75.0717
0.30116 | -75.1332
0.30149 | -75.1939
0.30182 | -75.2631
0.30215 | -75.3291
0.30248 | -75.4016
0.30281 | -75.4833
0.30314 | -75.5624
0.30346 | -75.6459
0.30379 | -75.7065
0.30412 | -75.7725
0.30445 | -75.846
0.30478 | -75.9336
0.30511 | -76.0059
0.30544 | -76.0819
0.30577 | -76.1683
0.3061 | -76.2285
0.30643 | -76.2795
0.30676 | -76.3153
0.30709 | -76.3573
0.30741 | -76.4018
0.30774 | -76.4441
0.30807 | -76.4855
0.3084 | -76.5165
0.30873 | -76.5474
0.30906 | -76.5688
0.30939 | -76.5799
0.30972 | -76.5621
0.31005 | -76.5727
0.31038 | -76.5689
0.31071 | -76.565
0.31104 | -76.5476
0.31136 | -76.5312
0.31169 | -76.5031
0.31202 | -76.4653
0.31235 | -76.4225
0.31268 | -76.3705
0.31301 | -76.3289
0.31334 | -76.2866
0.31367 | -76.2537
0.314 | -76.2133
0.31433 | -76.1734
0.31466 | -76.1435
0.31499 | -76.1221
0.31531 | -76.0924
0.31564 | -76.0727
0.31597 | -76.0441
0.3163 | -76.0192
0.31663 | -76.0194
0.31696 | -76.0147
0.31729 | -76.0044
0.31762 | -76.0072
0.31795 | -76.0106
0.31828 | -76.0264
0.31861 | -76.0332
0.31894 | -76.053
0.31926 | -76.0826
0.31959 | -76.1213
0.31992 | -76.1709
0.32025 | -76.2263
0.32058 | -76.2722
0.32091 | -76.3217
0.32124 | -76.3744
0.32157 | -76.4301
0.3219 | -76.4835
0.32223 | -76.5407
0.32255 | -76.6106
0.32288 | -76.6574
0.32321 | -76.6859
0.32354 | -76.7323
0.32387 | -76.7798
0.3242 | -76.7961
0.32453 | -76.8194
0.32486 | -76.8659
0.32519 | -76.8944
0.32552 | -76.9127
0.32585 | -76.9341
0.32618 | -76.9542
0.3265 | -76.9639
0.32683 | -76.9405
0.32716 | -76.9041
0.32749 | -76.8835
0.32782 | -76.8685
0.32815 | -76.8375
0.32848 | -76.7954
0.32881 | -76.753
0.32914 | -76.7237
0.32947 | -76.6961
0.3298 | -76.6452
0.33013 | -76.5827
0.33045 | -76.5158
0.33078 | -76.4566
0.33111 | -76.4038
0.33144 | -76.3484
0.33177 | -76.2901
0.3321 | -76.2355
0.33243 | -76.1774
0.33276 | -76.1366
0.33309 | -76.0769
0.33342 | -76.0398
0.33375 | -75.9921
0.33408 | -75.9465
0.3344 | -75.9088
0.33473 | -75.8712
0.33506 | -75.8451
0.33539 | -75.8264
0.33572 | -75.8151
0.33605 | -75.8052
0.33638 | -75.7972
0.33671 | -75.7972
0.33704 | -75.8023
0.33737 | -75.8114
0.3377 | -75.8271
0.33803 | -75.8474
0.33835 | -75.8759
0.33868 | -75.9031
0.33901 | -75.941
0.33934 | -75.9875
0.33967 | -76.0306
0.34 | -76.0795
0.34033 | -76.136
0.34066 | -76.1835
0.34099 | -76.2198
0.34132 | -76.2753
0.34164 | -76.3509
0.34197 | -76.4067
0.3423 | -76.486
0.34263 | -76.5495
0.34296 | -76.614
0.34329 | -76.6864
0.34362 | -76.7569
0.34395 | -76.8257
0.34428 | -76.8991
0.34461 | -76.9472
0.34494 | -76.9949
0.34527 | -77.059
0.34559 | -77.1261
0.34592 | -77.1936
0.34625 | -77.2526
0.34658 | -77.3033
0.34691 | -77.3674
0.34724 | -77.451
0.34757 | -77.5017
0.3479 | -77.5495
0.34823 | -77.6083
0.34856 | -77.6546
0.34889 | -77.6703
0.34921 | -77.6752
0.34954 | -77.6818
0.34987 | -77.6805
0.3502 | -77.6776
0.35053 | -77.6805
0.35086 | -77.7042
0.35119 | -77.6661
0.35152 | -77.6698
0.35185 | -77.644
0.35218 | -77.6168
0.35251 | -77.5973
0.35284 | -77.5709
0.35317 | -77.5378
0.35349 | -77.4989
0.35382 | -77.477
0.35415 | -77.4305
0.35448 | -77.3982
0.35481 | -77.3733
0.35514 | -77.3434
0.35547 | -77.322
0.3558 | -77.2964
0.35613 | -77.2757
0.35646 | -77.2645
0.35679 | -77.2596
0.35712 | -77.2503
0.35744 | -77.2465
0.35777 | -77.2404
0.3581 | -77.2275
0.35843 | -77.2228
0.35876 | -77.222
0.35909 | -77.2144
0.35942 | -77.2128
0.35975 | -77.2139
0.36008 | -77.2186
0.36041 | -77.2213
0.36073 | -77.2525
0.36106 | -77.268
0.36139 | -77.2831
0.36172 | -77.2985
0.36205 | -77.3143
0.36238 | -77.3314
0.36271 | -77.3568
0.36304 | -77.3839
0.36337 | -77.3971
0.3637 | -77.3939
0.36403 | -77.3878
0.36436 | -77.3944
0.36468 | -77.4128
0.36501 | -77.427
0.36534 | -77.4449
0.36567 | -77.4717
0.366 | -77.502
0.36633 | -77.5137
0.36666 | -77.5169
0.36699 | -77.5038
0.36732 | -77.5043
0.36765 | -77.5019
0.36798 | -77.5055
0.3683 | -77.5042
0.36863 | -77.4943
0.36896 | -77.4857
0.36929 | -77.4745
0.36962 | -77.4557
0.36995 | -77.4507
0.37028 | -77.415
0.37061 | -77.4096
0.37094 | -77.3871
0.37127 | -77.362
0.3716 | -77.3352
0.37193 | -77.3114
0.37226 | -77.29
0.37258 | -77.2662
0.37291 | -77.2299
0.37324 | -77.2043
0.37357 | -77.1834
0.3739 | -77.1642
0.37423 | -77.1456
0.37456 | -77.1311
0.37489 | -77.1193
0.37522 | -77.1096
0.37555 | -77.0964
0.37588 | -77.0909
0.37621 | -77.0887
0.37653 | -77.0753
0.37686 | -77.0626
0.37719 | -77.0666
0.37752 | -77.0771
0.37785 | -77.0939
0.37818 | -77.1157
0.37851 | -77.1404
0.37884 | -77.1731
0.37917 | -77.2104
0.3795 | -77.2498
0.37983 | -77.2995
0.38015 | -77.332
0.38048 | -77.3782
0.38081 | -77.41
0.38114 | -77.4412
0.38147 | -77.476
0.3818 | -77.4874
0.38213 | -77.5085
0.38246 | -77.5414
0.38279 | -77.5735
0.38312 | -77.6022
0.38345 | -77.6341
0.38377 | -77.655
0.3841 | -77.6682
0.38443 | -77.6817
0.38476 | -77.6917
0.38509 | -77.696
0.38542 | -77.7092
0.38575 | -77.7173
0.38608 | -77.7146
0.38641 | -77.7099
0.38674 | -77.7024
0.38707 | -77.6912
0.38739 | -77.6782
0.38772 | -77.6587
0.38805 | -77.6418
0.38838 | -77.6256
0.38871 | -77.6111
0.38904 | -77.5896
0.38937 | -77.5759
0.3897 | -77.5502
0.39003 | -77.5312
0.39036 | -77.5262
0.39069 | -77.5356
0.39102 | -77.5342
0.39134 | -77.5058
0.39167 | -77.4731
0.392 | -77.4507
0.39233 | -77.4308
0.39266 | -77.4137
0.39299 | -77.3997
0.39332 | -77.3906
0.39365 | -77.3875
0.39398 | -77.3985
0.39431 | -77.3995
0.39464 | -77.3963
0.39497 | -77.3881
0.3953 | -77.3863
0.39562 | -77.3985
0.39595 | -77.4065
0.39628 | -77.4117
0.39661 | -77.4308
0.39694 | -77.4579
0.39727 | -77.4812
0.3976 | -77.5075
0.39793 | -77.5361
0.39826 | -77.5522
0.39859 | -77.5825
0.39892 | -77.6089
0.39924 | -77.6347
0.39957 | -77.6661
0.3999 | -77.6996
0.40023 | -77.7354
0.40056 | -77.7792
0.40089 | -77.8158
0.40122 | -77.8397
0.40155 | -77.8562
0.40188 | -77.8747
0.40221 | -77.9027
0.40254 | -77.9323
0.40286 | -77.9572
0.40319 | -77.9894
0.40352 | -77.9992
0.40385 | -77.9951
0.40418 | -78.0217
0.40451 | -78.0522
0.40484 | -78.0624
0.40517 | -78.0598
0.4055 | -78.0396
0.40583 | -78.0293
0.40616 | -78.0209
0.40648 | -78.0095
0.40681 | -78.0002
0.40714 | -77.994
0.40747 | -77.9816
0.4078 | -77.9608
0.40813 | -77.9349
0.40846 | -77.9151
0.40879 | -77.9031
0.40912 | -77.8939
0.40945 | -77.8833
0.40978 | -77.87
0.41011 | -77.8723
0.41043 | -77.8637
0.41076 | -77.8417
0.41109 | -77.8114
0.41142 | -77.7904
0.41175 | -77.7787
0.41208 | -77.7566
0.41241 | -77.7414
0.41274 | -77.7283
0.41307 | -77.711
0.4134 | -77.6945
0.41373 | -77.6763
0.41406 | -77.6547
0.41439 | -77.6196
0.41471 | -77.604
0.41504 | -77.5948
0.41537 | -77.5859
0.4157 | -77.575
0.41603 | -77.5707
0.41636 | -77.5682
0.41669 | -77.5661
0.41702 | -77.5653
0.41735 | -77.563
0.41768 | -77.5813
0.41801 | -77.5864
0.41833 | -77.5874
0.41866 | -77.5958
0.41899 | -77.6039
0.41932 | -77.6068
0.41965 | -77.6205
0.41998 | -77.6316
0.42031 | -77.6456
0.42064 | -77.664
0.42097 | -77.6742
0.4213 | -77.6858
0.42163 | -77.6987
0.42196 | -77.7167
0.42228 | -77.7182
0.42261 | -77.7148
0.42294 | -77.7241
0.42327 | -77.7387
0.4236 | -77.7523
0.42393 | -77.7745
0.42426 | -77.7866
0.42459 | -77.7832
0.42492 | -77.7885
0.42525 | -77.7945
0.42557 | -77.7901
0.4259 | -77.7893
0.42623 | -77.7902
0.42656 | -77.7875
0.42689 | -77.7897
0.42722 | -77.7917
0.42755 | -77.7983
0.42788 | -77.8087
0.42821 | -77.8092
0.42854 | -77.8052
0.42887 | -77.8063
0.4292 | -77.805
0.42952 | -77.8041
0.42985 | -77.7925
0.43018 | -77.7693
0.43051 | -77.7545
0.43084 | -77.7511
0.43117 | -77.7621
0.4315 | -77.7609
0.43183 | -77.7552
0.43216 | -77.7331
0.43249 | -77.7224
0.43282 | -77.7159
0.43315 | -77.7207
0.43347 | -77.7304
0.4338 | -77.7169
0.43413 | -77.709
0.43446 | -77.6996
0.43479 | -77.6947
0.43512 | -77.7009
0.43545 | -77.6941
0.43578 | -77.6889
0.43611 | -77.6862
0.43644 | -77.6761
0.43677 | -77.6718
0.4371 | -77.6659
0.43742 | -77.6639
0.43775 | -77.6621
0.43808 | -77.6594
0.43841 | -77.659
0.43874 | -77.6628
0.43907 | -77.6556
0.4394 | -77.652
0.43973 | -77.6557
0.44006 | -77.6608
0.44039 | -77.667
0.44072 | -77.6741
0.44105 | -77.6837
0.44137 | -77.6981
0.4417 | -77.7063
0.44203 | -77.718
0.44236 | -77.7369
0.44269 | -77.7542
0.44302 | -77.7669
0.44335 | -77.7812
0.44368 | -77.794
0.44401 | -77.8129
0.44434 | -77.8382
0.44466 | -77.8734
0.44499 | -77.8941
0.44532 | -77.9157
0.44565 | -77.9351
0.44598 | -77.9548
0.44631 | -77.9732
0.44664 | -77.9927
0.44697 | -78.0095
0.4473 | -78.0297
0.44763 | -78.0537
0.44796 | -78.0823
0.44829 | -78.1043
0.44861 | -78.1266
0.44894 | -78.1496
0.44927 | -78.1777
0.4496 | -78.2029
0.44993 | -78.2215
0.45026 | -78.2391
0.45059 | -78.2536
0.45092 | -78.2807
0.45125 | -78.3105
0.45158 | -78.3301
0.45191 | -78.3437
0.45224 | -78.3479
0.45256 | -78.3638
0.45289 | -78.38
0.45322 | -78.4095
0.45355 | -78.4407
0.45388 | -78.4519
0.45421 | -78.4614
0.45454 | -78.4731
0.45487 | -78.4726
0.4552 | -78.4849
0.45553 | -78.4983
0.45586 | -78.5223
0.45619 | -78.541
0.45651 | -78.5601
0.45684 | -78.5778
0.45717 | -78.5916
0.4575 | -78.6023
0.45783 | -78.6139
0.45816 | -78.6205
0.45849 | -78.6308
0.45882 | -78.6389
0.45915 | -78.6565
0.45948 | -78.6662
0.45981 | -78.6783
0.46014 | -78.6993
0.46046 | -78.7074
0.46079 | -78.709
0.46112 | -78.7216
0.46145 | -78.7455
0.46178 | -78.7765
0.46211 | -78.804
0.46244 | -78.8243
0.46277 | -78.8355
0.4631 | -78.8555
0.46343 | -78.8732
0.46375 | -78.9079
0.46408 | -78.9258
0.46441 | -78.9432
0.46474 | -78.9629
0.46507 | -78.9763
0.4654 | -78.9901
0.46573 | -79.0139
0.46606 | -79.0353
0.46639 | -79.0561
0.46672 | -79.0778
0.46705 | -79.1074
0.46738 | -79.133
0.4677 | -79.1574
0.46803 | -79.1697
0.46836 | -79.1915
0.46869 | -79.227
0.46902 | -79.2452
0.46935 | -79.2505
0.46968 | -79.2576
0.47001 | -79.2771
0.47034 | -79.2955
0.47067 | -79.3065
0.471 | -79.3114
0.47133 | -79.3233
0.47165 | -79.3336
0.47198 | -79.3319
0.47231 | -79.3215
0.47264 | -79.3129
0.47297 | -79.3129
0.4733 | -79.3057
0.47363 | -79.3063
0.47396 | -79.31
0.47429 | -79.3006
0.47462 | -79.2715
0.47495 | -79.2628
0.47528 | -79.2475
0.4756 | -79.2334
0.47593 | -79.2254
0.47626 | -79.2181
0.47659 | -79.1991
0.47692 | -79.1808
0.47725 | -79.1623
0.47758 | -79.1224
0.47791 | -79.0674
0.47824 | -79.0274
0.47857 | -78.9973
0.4789 | -78.9721
0.47923 | -78.9515
0.47955 | -78.917
0.47988 | -78.8853
0.48021 | -78.8528
0.48054 | -78.8197
0.48087 | -78.788
0.4812 | -78.7613
0.48153 | -78.7238
0.48186 | -78.687
0.48219 | -78.6601
0.48252 | -78.6352
0.48285 | -78.6148
0.48317 | -78.6035
0.4835 | -78.5847
0.48383 | -78.5625
0.48416 | -78.5631
0.48449 | -78.5504
0.48482 | -78.5357
0.48515 | -78.5236
0.48548 | -78.5182
0.48581 | -78.5082
0.48614 | -78.5009
0.48647 | -78.4914
0.48679 | -78.4873
0.48712 | -78.4925
0.48745 | -78.498
0.48778 | -78.5044
0.48811 | -78.5198
0.48844 | -78.5267
0.48877 | -78.5337
0.4891 | -78.5442
0.48943 | -78.5576
0.48976 | -78.5735
0.49009 | -78.5887
0.49042 | -78.6098
0.49074 | -78.6327
0.49107 | -78.6488
0.4914 | -78.663
0.49173 | -78.683
0.49206 | -78.7003
0.49239 | -78.7298
0.49272 | -78.7477
0.49305 | -78.7692
0.49338 | -78.7908
0.49371 | -78.816
0.49404 | -78.8425
0.49436 | -78.8775
0.49469 | -78.9175
0.49502 | -78.9476
0.49535 | -78.9803
0.49568 | -79.0156
0.49601 | -79.0446
0.49634 | -79.0761
0.49667 | -79.1079
0.497 | -79.1293
0.49733 | -79.1371
0.49766 | -79.1567
0.49799 | -79.1851
0.49832 | -79.2114
0.49864 | -79.2361
0.49897 | -79.2653
0.4993 | -79.3014
0.49963 | -79.3265
0.49996 | -79.3482
0.50029 | -79.3588
0.50062 | -79.3734
0.50095 | -79.3901
0.50128 | -79.3972
0.50161 | -79.4107
0.50194 | -79.4253
0.50226 | -79.4501
0.50259 | -79.4693
0.50292 | -79.4802
0.50325 | -79.4879
0.50358 | -79.5047
0.50391 | -79.5266
0.50424 | -79.5371
0.50457 | -79.5407
0.5049 | -79.5458
0.50523 | -79.5512
0.50556 | -79.5607
0.50589 | -79.5716
0.50621 | -79.5789
0.50654 | -79.5872
0.50687 | -79.6025
0.5072 | -79.6147
0.50753 | -79.6172
0.50786 | -79.6248
0.50819 | -79.6229
0.50852 | -79.6254
0.50885 | -79.6274
0.50918 | -79.6251
0.50951 | -79.625
0.50984 | -79.6293
0.51016 | -79.6406
0.51049 | -79.6501
0.51082 | -79.6602
0.51115 | -79.6718
0.51148 | -79.6856
0.51181 | -79.7015
0.51214 | -79.7196
0.51247 | -79.7519
0.5128 | -79.7716
0.51313 | -79.8191
0.51345 | -79.8471
0.51378 | -79.8604
0.51411 | -79.8766
0.51444 | -79.8948
0.51477 | -79.9154
0.5151 | -79.9386
0.51543 | -79.9642
0.51576 | -79.9892
0.51609 | -80.013
0.51642 | -80.042
0.51675 | -80.0723
0.51708 | -80.1048
0.51741 | -80.1302
0.51773 | -80.1563
0.51806 | -80.174
0.51839 | -80.1938
0.51872 | -80.2182
0.51905 | -80.2396
0.51938 | -80.2662
0.51971 | -80.2858
0.52004 | -80.3055
0.52037 | -80.3309
0.5207 | -80.3597
0.52102 | -80.3921
0.52135 | -80.4158
0.52168 | -80.4379
0.52201 | -80.4314
0.52234 | -80.4524
0.52267 | -80.4745
0.523 | -80.4931
0.52333 | -80.5078
0.52366 | -80.5267
0.52399 | -80.5419
0.52432 | -80.5517
0.52465 | -80.5473
0.52497 | -80.5519
0.5253 | -80.5538
0.52563 | -80.5571
0.52596 | -80.5663
0.52629 | -80.5714
0.52662 | -80.5708
0.52695 | -80.5781
0.52728 | -80.5872
0.52761 | -80.5918
0.52794 | -80.5997
0.52827 | -80.5938
0.5286 | -80.6
0.52892 | -80.6001
0.52925 | -80.5924
0.52958 | -80.5789
0.52991 | -80.5634
0.53024 | -80.542
0.53057 | -80.5247
0.5309 | -80.509
0.53123 | -80.4913
0.53156 | -80.4713
0.53189 | -80.4497
0.53222 | -80.4278
0.53255 | -80.4112
0.53287 | -80.3891
0.5332 | -80.3678
0.53353 | -80.3542
0.53386 | -80.3588
0.53419 | -80.3496
0.53452 | -80.3315
0.53485 | -80.3081
0.53518 | -80.2863
0.53551 | -80.2701
0.53584 | -80.2483
0.53617 | -80.2306
0.5365 | -80.2118
0.53682 | -80.1997
0.53715 | -80.182
0.53748 | -80.1608
0.53781 | -80.1503
0.53814 | -80.1368
0.53847 | -80.1211
0.5388 | -80.1059
0.53913 | -80.0921
0.53946 | -80.0802
0.53979 | -80.067
0.54012 | -80.0444
0.54044 | -80.0233
0.54077 | -80.0182
0.5411 | -80.0225
0.54143 | -80.0281
0.54176 | -80.028
0.54209 | -80.0223
0.54242 | -80.0182
0.54275 | -80.0184
0.54308 | -80.0204
0.54341 | -80.0181
0.54374 | -80.0195
0.54407 | -80.0237
0.54439 | -80.0278
0.54472 | -80.0344
0.54505 | -80.0405
0.54538 | -80.0522
0.54571 | -80.0661
0.54604 | -80.0789
0.54637 | -80.0894
0.5467 | -80.1029
0.54703 | -80.1185
0.54736 | -80.1285
0.54769 | -80.1418
0.54802 | -80.1586
0.54834 | -80.1771
0.54867 | -80.1941
0.549 | -80.2071
0.54933 | -80.224
0.54966 | -80.2462
0.54999 | -80.2787
0.55032 | -80.3006
0.55065 | -80.3238
0.55098 | -80.3421
0.55131 | -80.3653
0.55163 | -80.3958
0.55196 | -80.4214
0.55229 | -80.4473
0.55262 | -80.4695
0.55295 | -80.4879
0.55328 | -80.5051
0.55361 | -80.5102
0.55394 | -80.5113
0.55427 | -80.5211
0.5546 | -80.5432
0.55493 | -80.5654
0.55526 | -80.5821
0.55558 | -80.5981
0.55591 | -80.6214
0.55624 | -80.6389
0.55657 | -80.6531
0.5569 | -80.6642
0.55723 | -80.6633
0.55756 | -80.6706
0.55789 | -80.6779
0.55822 | -80.6745
0.55855 | -80.6759
0.55888 | -80.6733
0.5592 | -80.6742
0.55953 | -80.6771
0.55986 | -80.6854
0.56019 | -80.6642
0.56052 | -80.6749
0.56085 | -80.6732
0.56118 | -80.6639
0.56151 | -80.655
0.56184 | -80.6509
0.56217 | -80.6426
0.5625 | -80.6389
0.56283 | -80.6373
0.56315 | -80.6438
0.56348 | -80.6385
0.56381 | -80.6315
0.56414 | -80.6234
0.56447 | -80.6074
0.5648 | -80.5957
0.56513 | -80.585
0.56546 | -80.5784
0.56579 | -80.5716
0.56612 | -80.5606
0.56645 | -80.545
0.56678 | -80.5324
0.5671 | -80.523
0.56743 | -80.5137
0.56776 | -80.5078
0.56809 | -80.5072
0.56842 | -80.5002
0.56875 | -80.4989
0.56908 | -80.4933
0.56941 | -80.4824
0.56974 | -80.4917
0.57007 | -80.4903
0.5704 | -80.4881
0.57073 | -80.4913
0.57105 | -80.4912
0.57138 | -80.4858
0.57171 | -80.4772
0.57204 | -80.4783
0.57237 | -80.4744
0.5727 | -80.4749
0.57303 | -80.4715
0.57336 | -80.4748
0.57369 | -80.4799
0.57402 | -80.486
0.57435 | -80.4924
0.57468 | -80.4984
0.575 | -80.5077
0.57533 | -80.5113
0.57566 | -80.5099
0.57599 | -80.5145
0.57632 | -80.5255
0.57665 | -80.5368
0.57698 | -80.5457
0.57731 | -80.5554
0.57764 | -80.564
0.57797 | -80.5739
0.5783 | -80.5729
0.57862 | -80.554
0.57895 | -80.5431
0.57928 | -80.5442
0.57961 | -80.5554
0.57994 | -80.5668
0.58027 | -80.5675
0.5806 | -80.5682
0.58093 | -80.5673
0.58126 | -80.562
0.58159 | -80.5603
0.58192 | -80.5536
0.58225 | -80.5469
0.58257 | -80.5425
0.5829 | -80.5362
0.58323 | -80.532
0.58356 | -80.5274
0.58389 | -80.5167
0.58422 | -80.4991
0.58455 | -80.4749
0.58488 | -80.462
0.58521 | -80.4531
0.58554 | -80.4354
0.58587 | -80.4228
0.5862 | -80.42
0.58652 | -80.4117
0.58685 | -80.3938
0.58718 | -80.3798
0.58751 | -80.3648
0.58784 | -80.3474
0.58817 | -80.3289
0.5885 | -80.2995
0.58883 | -80.2801
0.58916 | -80.2671
0.58949 | -80.251
0.58981 | -80.2336
0.59014 | -80.2161
0.59047 | -80.2036
0.5908 | -80.1946
0.59113 | -80.183
0.59146 | -80.1732
0.59179 | -80.1615
0.59212 | -80.1509
0.59245 | -80.143
0.59278 | -80.1348
0.59311 | -80.1361
0.59344 | -80.1389
0.59376 | -80.1411
0.59409 | -80.1521
0.59442 | -80.1671
0.59475 | -80.1743
0.59508 | -80.1784
0.59541 | -80.1806
0.59574 | -80.1867
0.59607 | -80.1966
0.5964 | -80.2069
0.59673 | -80.2198
0.59706 | -80.2362
0.59739 | -80.257
0.59771 | -80.2848
0.59804 | -80.3322
0.59837 | -80.3606
0.5987 | -80.3894
0.59903 | -80.4216
0.59936 | -80.4501
0.59969 | -80.4863
0.60002 | -80.5223
0.60035 | -80.5564
0.60068 | -80.5913
0.60101 | -80.638
0.60133 | -80.6854
0.60166 | -80.7307
0.60199 | -80.7784
0.60232 | -80.8287
0.60265 | -80.8783
0.60298 | -80.9318
0.60331 | -80.987
0.60364 | -81.0448
0.60397 | -81.1046
0.6043 | -81.1645
0.60463 | -81.2257
0.60496 | -81.282
0.60528 | -81.345
0.60561 | -81.4163
0.60594 | -81.4854
0.60627 | -81.5557
0.6066 | -81.6217
0.60693 | -81.6898
0.60726 | -81.7625
0.60759 | -81.8315
0.60792 | -81.9079
0.60825 | -81.9783
0.60858 | -82.0493
0.60891 | -82.1254
0.60923 | -82.2035
0.60956 | -82.2784
0.60989 | -82.359
0.61022 | -82.433
0.61055 | -82.5076
0.61088 | -82.589
0.61121 | -82.6649
0.61154 | -82.7259
0.61187 | -82.7784
0.6122 | -82.8316
0.61253 | -82.8939
0.61286 | -82.9685
0.61318 | -83.0493
0.61351 | -83.1248
0.61384 | -83.1758
0.61417 | -83.2173
0.6145 | -83.2578
0.61483 | -83.3056
0.61516 | -83.3416
0.61549 | -83.3712
0.61582 | -83.4103
0.61615 | -83.4457
0.61647 | -83.459
0.6168 | -83.4639
0.61713 | -83.4321
0.61746 | -83.4596
0.61779 | -83.4814
0.61812 | -83.4934
0.61845 | -83.4909
0.61878 | -83.473
0.61911 | -83.4531
0.61944 | -83.4357
0.61977 | -83.4163
0.6201 | -83.3905
0.62043 | -83.3818
0.62075 | -83.4023
0.62108 | -83.3844
0.62141 | -83.3351
0.62174 | -83.2889
0.62207 | -83.2433
0.6224 | -83.1976
0.62273 | -83.1498
0.62306 | -83.0981
0.62339 | -83.0436
0.62372 | -82.9838
0.62405 | -82.9102
0.62438 | -82.8388
0.6247 | -82.7685
0.62503 | -82.7008
0.62536 | -82.6399
0.62569 | -82.5795
0.62602 | -82.5117
0.62635 | -82.4538
0.62668 | -82.3807
0.62701 | -82.3303
0.62734 | -82.285
0.62767 | -82.2254
0.62799 | -82.1617
0.62832 | -82.1001
0.62865 | -82.0352
0.62898 | -81.9714
0.62931 | -81.9116
0.62964 | -81.8499
0.62997 | -81.7861
0.6303 | -81.7263
0.63063 | -81.671
0.63096 | -81.6168
0.63129 | -81.5725
0.63162 | -81.5326
0.63194 | -81.4833
0.63227 | -81.4362
0.6326 | -81.3835
0.63293 | -81.3313
0.63326 | -81.2851
0.63359 | -81.2402
0.63392 | -81.1922
0.63425 | -81.1445
0.63458 | -81.0974
0.63491 | -81.0591
0.63524 | -81.0252
0.63557 | -80.9924
0.63589 | -80.9627
0.63622 | -80.9515
0.63655 | -80.9166
0.63688 | -80.8818
0.63721 | -80.853
0.63754 | -80.8321
0.63787 | -80.8135
0.6382 | -80.7922
0.63853 | -80.7636
0.63886 | -80.7403
0.63919 | -80.724
0.63951 | -80.7074
0.63984 | -80.6928
0.64017 | -80.6803
0.6405 | -80.6706
0.64083 | -80.6638
0.64116 | -80.6481
0.64149 | -80.6226
0.64182 | -80.6152
0.64215 | -80.6046
0.64248 | -80.5931
0.64281 | -80.591
0.64314 | -80.5914
0.64346 | -80.5896
0.64379 | -80.587
0.64412 | -80.5885
0.64445 | -80.5941
0.64478 | -80.5957
0.64511 | -80.598
0.64544 | -80.5774
0.64577 | -80.5667
0.6461 | -80.5633
0.64643 | -80.5643
0.64676 | -80.5667
0.64709 | -80.5628
0.64741 | -80.5575
0.64774 | -80.5535
0.64807 | -80.5539
0.6484 | -80.5514
0.64873 | -80.5476
0.64906 | -80.5522
0.64939 | -80.5481
0.64972 | -80.5367
0.65005 | -80.5269
0.65038 | -80.519
0.65071 | -80.5161
0.65104 | -80.5127
0.65136 | -80.5049
0.65169 | -80.4958
0.65202 | -80.4881
0.65235 | -80.4796
0.65268 | -80.4728
0.65301 | -80.4652
0.65334 | -80.4543
0.65367 | -80.4435
0.654 | -80.4332
0.65433 | -80.4221
0.65465 | -80.4075
0.65498 | -80.3887
0.65531 | -80.3766
0.65564 | -80.3669
0.65597 | -80.3567
0.6563 | -80.3507
0.65663 | -80.3482
0.65696 | -80.3429
0.65729 | -80.3334
0.65762 | -80.3225
0.65795 | -80.3112
0.65828 | -80.2972
0.6586 | -80.2917
0.65893 | -80.2951
0.65926 | -80.295
0.65959 | -80.2889
0.65992 | -80.282
0.66025 | -80.2776
0.66058 | -80.2704
0.66091 | -80.2655
0.66124 | -80.2685
0.66157 | -80.2723
0.6619 | -80.2761
0.66222 | -80.2759
0.66256 | -80.2732
0.66288 | -80.2679
0.66321 | -80.2671
0.66354 | -80.2721
0.66387 | -80.2821
0.6642 | -80.2908
0.66453 | -80.3165
0.66486 | -80.3267
0.66519 | -80.3386
0.66552 | -80.3535
0.66585 | -80.3688
0.66617 | -80.3813
0.6665 | -80.3944
0.66683 | -80.4158
0.66716 | -80.4406
0.66749 | -80.4627
0.66782 | -80.4829
0.66815 | -80.5037
0.66848 | -80.5271
0.66881 | -80.5543
0.66914 | -80.5864
0.66947 | -80.6236
0.6698 | -80.6601
0.67012 | -80.6909
0.67045 | -80.7207
0.67078 | -80.7505
0.67111 | -80.7812
0.67144 | -80.8148
0.67177 | -80.8445
0.6721 | -80.8748
0.67243 | -80.9105
0.67276 | -80.9473
0.67309 | -80.9841
0.67342 | -81.0227
0.67375 | -81.0626
0.67407 | -81.1045
0.6744 | -81.1442
0.67473 | -81.1872
0.67506 | -81.2299
0.67539 | -81.2712
0.67572 | -81.3126
0.67605 | -81.3566
0.67638 | -81.3976
0.67671 | -81.4353
0.67704 | -81.4765
0.67737 | -81.5185
0.6777 | -81.558
0.67802 | -81.6029
0.67835 | -81.6373
0.67868 | -81.6787
0.67901 | -81.7189
0.67934 | -81.7528
0.67967 | -81.7892
0.68 | -81.8284
0.68033 | -81.8682
0.68066 | -81.9066
0.68099 | -81.9362
0.68132 | -81.9562
0.68164 | -81.9756
0.68197 | -82.0029
0.6823 | -82.0304
0.68263 | -82.063
0.68296 | -82.0965
0.68329 | -82.1222
0.68362 | -82.1291
0.68395 | -82.1542
0.68428 | -82.1743
0.68461 | -82.1924
0.68494 | -82.2096
0.68527 | -82.2371
0.68559 | -82.259
0.68592 | -82.2705
0.68625 | -82.2864
0.68658 | -82.3002
0.68691 | -82.3131
0.68724 | -82.3216
0.68757 | -82.3192
0.6879 | -82.3292
0.68823 | -82.3595
0.68856 | -82.3612
0.68889 | -82.3678
0.68922 | -82.3746
0.68954 | -82.3737
0.68987 | -82.3699
0.6902 | -82.366
0.69053 | -82.3637
0.69086 | -82.3583
0.69119 | -82.3482
0.69152 | -82.3377
0.69185 | -82.3256
0.69218 | -82.3292
0.69251 | -82.3379
0.69283 | -82.3384
0.69316 | -82.3272
0.69349 | -82.319
0.69382 | -82.3052
0.69415 | -82.292
0.69448 | -82.2818
0.69481 | -82.2739
0.69514 | -82.2559
0.69547 | -82.2354
0.6958 | -82.2228
0.69613 | -82.2134
0.69646 | -82.201
0.69678 | -82.1925
0.69711 | -82.1898
0.69744 | -82.1828
0.69777 | -82.1734
0.6981 | -82.1621
0.69843 | -82.1508
0.69876 | -82.138
0.69909 | -82.1135
0.69942 | -82.0928
0.69975 | -82.0823
0.70008 | -82.0686
0.70041 | -82.0489
0.70073 | -82.0294
0.70106 | -82.0203
0.70139 | -82.0115
0.70172 | -82.0039
0.70205 | -81.9998
0.70238 | -81.9978
0.70271 | -81.9832
0.70304 | -81.9799
0.70337 | -81.9705
0.7037 | -81.958
0.70403 | -81.9532
0.70435 | -81.9342
0.70468 | -81.9235
0.70501 | -81.9213
0.70534 | -81.9172
0.70567 | -81.9132
0.706 | -81.9079
0.70633 | -81.903
0.70666 | -81.9005
0.70699 | -81.8962
0.70732 | -81.8911
0.70765 | -81.8821
0.70798 | -81.8756
0.7083 | -81.8729
0.70863 | -81.8733
0.70896 | -81.8725
0.70929 | -81.8672
0.70962 | -81.8645
0.70995 | -81.861
0.71028 | -81.8578
0.71061 | -81.8559
0.71094 | -81.8647
0.71127 | -81.8694
0.7116 | -81.8701
0.71193 | -81.8577
0.71225 | -81.8574
0.71258 | -81.8553
0.71291 | -81.853
0.71324 | -81.8504
0.71357 | -81.8462
0.7139 | -81.8424
0.71423 | -81.8382
0.71456 | -81.8337
0.71489 | -81.8321
0.71522 | -81.8316
0.71555 | -81.8314
0.71588 | -81.8281
0.7162 | -81.8245
0.71653 | -81.82
0.71686 | -81.8166
0.71719 | -81.8147
0.71752 | -81.8135
0.71785 | -81.8073
0.71818 | -81.7997
0.71851 | -81.7997
0.71884 | -81.8041
0.71917 | -81.8038
0.71949 | -81.8018
0.71983 | -81.7998
0.72015 | -81.7973
0.72048 | -81.798
0.72081 | -81.8006
0.72114 | -81.7992
0.72147 | -81.8059
0.7218 | -81.806
0.72213 | -81.8027
0.72246 | -81.804
0.72279 | -81.8026
0.72312 | -81.8024
0.72345 | -81.8075
0.72377 | -81.8042
0.7241 | -81.8025
0.72443 | -81.8046
0.72476 | -81.8052
0.72509 | -81.8115
0.72542 | -81.8202
0.72575 | -81.8293
0.72608 | -81.8392
0.72641 | -81.8483
0.72674 | -81.8546
0.72707 | -81.8611
0.7274 | -81.8701
0.72772 | -81.8845
0.72805 | -81.8976
0.72838 | -81.9139
0.72871 | -81.9328
0.72904 | -81.9485
0.72937 | -81.9658
0.7297 | -81.9843
0.73003 | -82.0045
0.73036 | -82.0267
0.73069 | -82.0521
0.73101 | -82.0797
0.73134 | -82.1036
0.73167 | -82.1155
0.732 | -82.1191
0.73233 | -82.1288
0.73266 | -82.1478
0.73299 | -82.1735
0.73332 | -82.1927
0.73365 | -82.2094
0.73398 | -82.2299
0.73431 | -82.2474
0.73464 | -82.2636
0.73496 | -82.2855
0.73529 | -82.3084
0.73562 | -82.3307
0.73595 | -82.3488
0.73628 | -82.3685
0.73661 | -82.388
0.73694 | -82.4087
0.73727 | -82.4289
0.7376 | -82.4584
0.73793 | -82.4868
0.73826 | -82.5048
0.73859 | -82.5252
0.73891 | -82.5461
0.73924 | -82.5654
0.73957 | -82.5829
0.7399 | -82.6001
0.74023 | -82.6192
0.74056 | -82.6317
0.74089 | -82.6487
0.74122 | -82.6682
0.74155 | -82.6892
0.74188 | -82.7101
0.74221 | -82.7296
0.74253 | -82.7495
0.74286 | -82.7717
0.74319 | -82.7915
0.74352 | -82.8076
0.74385 | -82.8235
0.74418 | -82.8393
0.74451 | -82.8554
0.74484 | -82.8706
0.74517 | -82.8861
0.7455 | -82.9008
0.74583 | -82.9216
0.74616 | -82.9415
0.74648 | -82.9568
0.74681 | -82.9642
0.74714 | -82.9726
0.74747 | -82.9815
0.7478 | -82.9836
0.74813 | -82.9928
0.74846 | -83.0052
0.74879 | -83.016
0.74912 | -83.0279
0.74945 | -83.0451
0.74978 | -83.0733
0.75011 | -83.0872
0.75043 | -83.1075
0.75076 | -83.1138
0.75109 | -83.1307
0.75142 | -83.153
0.75175 | -83.1758
0.75208 | -83.1866
0.75241 | -83.1967
0.75274 | -83.2041
0.75307 | -83.2125
0.7534 | -83.2202
0.75373 | -83.237
0.75406 | -83.2512
0.75438 | -83.2599
0.75471 | -83.2725
0.75504 | -83.2894
0.75537 | -83.2988
0.7557 | -83.3256
0.75603 | -83.341
0.75636 | -83.3475
0.75669 | -83.3546
0.75702 | -83.3591
0.75735 | -83.3706
0.75767 | -83.388
0.75801 | -83.4011
0.75833 | -83.4132
0.75866 | -83.4267
0.75899 | -83.4463
0.75932 | -83.4625
0.75965 | -83.4782
0.75998 | -83.4886
0.76031 | -83.4993
0.76064 | -83.5107
0.76097 | -83.5269
0.7613 | -83.5428
0.76162 | -83.5578
0.76195 | -83.5705
0.76228 | -83.5886
0.76261 | -83.6041
0.76294 | -83.6159
0.76327 | -83.6198
0.7636 | -83.6185
0.76393 | -83.6212
0.76426 | -83.628
0.76459 | -83.6369
0.76492 | -83.639
0.76525 | -83.6392
0.76558 | -83.6375
0.7659 | -83.6302
0.76623 | -83.6285
0.76656 | -83.6296
0.76689 | -83.6267
0.76722 | -83.6263
0.76755 | -83.6264
0.76788 | -83.6276
0.76821 | -83.6295
0.76854 | -83.6306
0.76887 | -83.628
0.76919 | -83.6024
0.76952 | -83.593
0.76985 | -83.5898
0.77018 | -83.5889
0.77051 | -83.5879
0.77084 | -83.5785
0.77117 | -83.569
0.7715 | -83.5569
0.77183 | -83.5466
0.77216 | -83.5379
0.77249 | -83.5278
0.77282 | -83.5144
0.77314 | -83.4994
0.77347 | -83.4878
0.7738 | -83.4713
0.77413 | -83.4567
0.77446 | -83.439
0.77479 | -83.4271
0.77512 | -83.4118
0.77545 | -83.3991
0.77578 | -83.3867
0.77611 | -83.3711
0.77644 | -83.3634
0.77677 | -83.3516
0.77709 | -83.3314
0.77742 | -83.3101
0.77775 | -83.2886
0.77808 | -83.2672
0.77841 | -83.2448
0.77874 | -83.2238
0.77907 | -83.2042
0.7794 | -83.1843
0.77973 | -83.1675
0.78006 | -83.1528
0.78039 | -83.1286
0.78072 | -83.1077
0.78104 | -83.0911
0.78137 | -83.0709
0.7817 | -83.0488
0.78203 | -83.0281
0.78236 | -83.008
0.78269 | -82.9883
0.78302 | -82.9687
0.78335 | -82.949
0.78368 | -82.9307
0.78401 | -82.9151
0.78434 | -82.8964
0.78466 | -82.8793
0.78499 | -82.8685
0.78532 | -82.8675
0.78565 | -82.8576
0.78598 | -82.8419
0.78631 | -82.8269
0.78664 | -82.8113
0.78697 | -82.7965
0.7873 | -82.7841
0.78763 | -82.7732
0.78796 | -82.7693
0.78829 | -82.7627
0.78861 | -82.7558
0.78894 | -82.7452
0.78927 | -82.7323
0.7896 | -82.7301
0.78993 | -82.7256
0.79026 | -82.7227
0.79059 | -82.7168
0.79092 | -82.7142
0.79125 | -82.7132
0.79158 | -82.7111
0.79191 | -82.7069
0.79224 | -82.7069
0.79256 | -82.7062
0.79289 | -82.7094
0.79322 | -82.7147
0.79355 | -82.7221
0.79388 | -82.7288
0.79421 | -82.7371
0.79454 | -82.7475
0.79487 | -82.7574
0.7952 | -82.7656
0.79553 | -82.7784
0.79585 | -82.7905
0.79619 | -82.8042
0.79651 | -82.8176
0.79684 | -82.8319
0.79717 | -82.8462
0.7975 | -82.8737
0.79783 | -82.8898
0.79816 | -82.9079
0.79849 | -82.9193
0.79882 | -82.9311
0.79915 | -82.952
0.79948 | -82.9728
0.7998 | -82.9919
0.80013 | -83.0168
0.80046 | -83.0394
0.80079 | -83.0608
0.80112 | -83.088
0.80145 | -83.1135
0.80178 | -83.14
0.80211 | -83.1726
0.80244 | -83.1978
0.80277 | -83.2202
0.8031 | -83.2438
0.80343 | -83.2677
0.80375 | -83.293
0.80408 | -83.3204
0.80441 | -83.3458
0.80474 | -83.3718
0.80507 | -83.3981
0.8054 | -83.4277
0.80573 | -83.4526
0.80606 | -83.4766
0.80639 | -83.5022
0.80672 | -83.527
0.80705 | -83.5519
0.80737 | -83.5749
0.8077 | -83.6042
0.80803 | -83.6339
0.80836 | -83.6585
0.80869 | -83.681
0.80902 | -83.7117
0.80935 | -83.7384
0.80968 | -83.7603
0.81001 | -83.7819
0.81034 | -83.8028
0.81067 | -83.8268
0.811 | -83.8513
0.81132 | -83.876
0.81165 | -83.8977
0.81198 | -83.9249
0.81231 | -83.9414
0.81264 | -83.9517
0.81297 | -83.9684
0.8133 | -83.9853
0.81363 | -83.9902
0.81396 | -84.0037
0.81429 | -84.0211
0.81462 | -84.0383
0.81495 | -84.0573
0.81527 | -84.071
0.8156 | -84.0828
0.81593 | -84.0964
0.81626 | -84.1068
0.81659 | -84.1137
0.81692 | -84.1299
0.81725 | -84.1429
0.81758 | -84.1574
0.81791 | -84.1713
0.81824 | -84.1825
0.81857 | -84.189
0.8189 | -84.2015
0.81922 | -84.2339
0.81955 | -84.2562
0.81988 | -84.2631
0.82021 | -84.2727
0.82054 | -84.2792
0.82087 | -84.276
0.8212 | -84.2834
0.82153 | -84.2994
0.82186 | -84.3096
0.82219 | -84.3166
0.82252 | -84.3261
0.82285 | -84.3348
0.82317 | -84.3359
0.8235 | -84.3345
0.82383 | -84.3346
0.82416 | -84.3344
0.82449 | -84.3278
0.82482 | -84.3301
0.82515 | -84.3332
0.82548 | -84.336
0.82581 | -84.3313
0.82614 | -84.3139
0.82647 | -84.3054
0.82679 | -84.3001
0.82712 | -84.292
0.82745 | -84.2837
0.82778 | -84.2771
0.82811 | -84.2625
0.82844 | -84.2403
0.82877 | -84.2224
0.8291 | -84.2178
0.82943 | -84.205
0.82976 | -84.1935
0.83009 | -84.187
0.83042 | -84.1721
0.83074 | -84.1569
0.83107 | -84.1458
0.8314 | -84.1315
0.83173 | -84.1155
0.83206 | -84.1004
0.83239 | -84.0854
0.83272 | -84.0683
0.83305 | -84.0489
0.83338 | -84.0365
0.83371 | -84.0282
0.83403 | -84.0152
0.83437 | -84
0.83469 | -83.9829
0.83502 | -83.9666
0.83535 | -83.952
0.83568 | -83.9275
0.83601 | -83.9057
0.83634 | -83.886
0.83667 | -83.8695
0.837 | -83.8542
0.83733 | -83.8429
0.83766 | -83.8303
0.83798 | -83.8138
0.83831 | -83.799
0.83864 | -83.7841
0.83897 | -83.7684
0.8393 | -83.7506
0.83963 | -83.7348
0.83996 | -83.7213
0.84029 | -83.7056
0.84062 | -83.6861
0.84095 | -83.6683
0.84128 | -83.6572
0.84161 | -83.6474
0.84193 | -83.6339
0.84226 | -83.6182
0.84259 | -83.6018
0.84292 | -83.5863
0.84325 | -83.5733
0.84358 | -83.563
0.84391 | -83.5551
0.84424 | -83.5446
0.84457 | -83.5343
0.8449 | -83.5328
0.84523 | -83.5316
0.84555 | -83.5273
0.84588 | -83.5209
0.84621 | -83.5136
0.84654 | -83.5098
0.84687 | -83.5014
0.8472 | -83.4951
0.84753 | -83.4889
0.84786 | -83.4774
0.84819 | -83.4554
0.84852 | -83.4388
0.84885 | -83.4279
0.84918 | -83.4217
0.8495 | -83.4211
0.84983 | -83.4182
0.85016 | -83.415
0.85049 | -83.4073
0.85082 | -83.3995
0.85115 | -83.3924
0.85148 | -83.3867
0.85181 | -83.3851
0.85214 | -83.3833
0.85247 | -83.3819
0.8528 | -83.3823
0.85313 | -83.3801
0.85345 | -83.3773
0.85378 | -83.3784
0.85411 | -83.376
0.85444 | -83.376
0.85477 | -83.3632
0.8551 | -83.3515
0.85543 | -83.3481
0.85576 | -83.3481
0.85609 | -83.3465
0.85642 | -83.3459
0.85675 | -83.3528
0.85708 | -83.3575
0.8574 | -83.3549
0.85773 | -83.3527
0.85806 | -83.355
0.85839 | -83.3519
0.85872 | -83.3512
0.85905 | -83.35
0.85938 | -83.3489
0.85971 | -83.3488
0.86004 | -83.3478
0.86037 | -83.3473
0.8607 | -83.3486
0.86103 | -83.3487
0.86135 | -83.3473
0.86168 | -83.3415
0.86201 | -83.3364
0.86234 | -83.3274
0.86267 | -83.3225
0.863 | -83.3206
0.86333 | -83.3194
0.86366 | -83.3155
0.86399 | -83.3103
0.86432 | -83.3042
0.86464 | -83.301
0.86497 | -83.2981
0.8653 | -83.2958
0.86563 | -83.2918
0.86596 | -83.2885
0.86629 | -83.2846
0.86662 | -83.2756
0.86695 | -83.2691
0.86728 | -83.2637
0.86761 | -83.2601
0.86794 | -83.2561
0.86827 | -83.246
0.8686 | -83.2378
0.86892 | -83.2314
0.86925 | -83.226
0.86958 | -83.217
0.86991 | -83.211
0.87024 | -83.2033
0.87057 | -83.1993
0.8709 | -83.1927
0.87123 | -83.1904
0.87156 | -83.189
0.87189 | -83.1845
0.87221 | -83.1768
0.87255 | -83.1784
0.87287 | -83.1726
0.8732 | -83.1592
0.87353 | -83.1477
0.87386 | -83.1382
0.87419 | -83.1314
0.87452 | -83.128
0.87485 | -83.1246
0.87518 | -83.1203
0.87551 | -83.113
0.87584 | -83.1061
0.87616 | -83.1004
0.87649 | -83.0949
0.87682 | -83.089
0.87715 | -83.0846
0.87748 | -83.0807
0.87781 | -83.0779
0.87814 | -83.0792
0.87847 | -83.0826
0.8788 | -83.0818
0.87913 | -83.0813
0.87946 | -83.082
0.87979 | -83.08
0.88011 | -83.0842
0.88044 | -83.0893
0.88077 | -83.0882
0.8811 | -83.0868
0.88143 | -83.0869
0.88176 | -83.0864
0.88209 | -83.0827
0.88242 | -83.0856
0.88275 | -83.0885
0.88308 | -83.0963
0.88341 | -83.1
0.88374 | -83.105
0.88406 | -83.1104
0.88439 | -83.1192
0.88472 | -83.1205
0.88505 | -83.1117
0.88538 | -83.1108
0.88571 | -83.1162
0.88604 | -83.1245
0.88637 | -83.1354
0.8867 | -83.142
0.88703 | -83.1492
0.88736 | -83.1543
0.88768 | -83.1619
0.88801 | -83.1692
0.88834 | -83.1787
0.88867 | -83.1931
0.889 | -83.2046
0.88933 | -83.2112
0.88966 | -83.2206
0.88999 | -83.228
0.89032 | -83.2352
0.89065 | -83.2459
0.89098 | -83.2566
0.89131 | -83.2635
0.89163 | -83.2688
0.89196 | -83.2788
0.89229 | -83.2872
0.89262 | -83.2969
0.89295 | -83.3049
0.89328 | -83.3116
0.89361 | -83.3236
0.89394 | -83.3323
0.89427 | -83.3414
0.8946 | -83.3485
0.89493 | -83.3569
0.89526 | -83.363
0.89558 | -83.3691
0.89591 | -83.378
0.89624 | -83.3889
0.89657 | -83.4038
0.8969 | -83.4133
0.89723 | -83.4163
0.89756 | -83.4197
0.89789 | -83.426
0.89822 | -83.4304
0.89855 | -83.4324
0.89888 | -83.4359
0.89921 | -83.4415
0.89953 | -83.4488
0.89986 | -83.4566
0.90019 | -83.4643
0.90052 | -83.4701
0.90085 | -83.4749
0.90118 | -83.4808
0.90151 | -83.4925
0.90184 | -83.4947
0.90217 | -83.5011
0.9025 | -83.5068
0.90282 | -83.5126
0.90315 | -83.519
0.90348 | -83.5165
0.90381 | -83.5132
0.90414 | -83.5147
0.90447 | -83.5188
0.9048 | -83.5211
0.90513 | -83.5214
0.90546 | -83.5213
0.90579 | -83.5236
0.90612 | -83.5284
0.90645 | -83.5279
0.90677 | -83.5293
0.9071 | -83.5309
0.90743 | -83.5323
0.90776 | -83.5335
0.90809 | -83.5368
0.90842 | -83.547
0.90875 | -83.5469
0.90908 | -83.5454
0.90941 | -83.5527
0.90974 | -83.5712
0.91007 | -83.5837
0.91039 | -83.5884
0.91073 | -83.5919
0.91105 | -83.5943
0.91138 | -83.598
0.91171 | -83.5979
0.91204 | -83.5925
0.91237 | -83.5971
0.9127 | -83.5976
0.91303 | -83.5996
0.91336 | -83.6016
0.91369 | -83.6012
0.91402 | -83.6022
0.91434 | -83.6007
0.91467 | -83.6005
0.915 | -83.6008
0.91533 | -83.6021
0.91566 | -83.6036
0.91599 | -83.6094
0.91632 | -83.6168
0.91665 | -83.6173
0.91698 | -83.6258
0.91731 | -83.6344
0.91764 | -83.6377
0.91797 | -83.6403
0.91829 | -83.643
0.91862 | -83.6431
0.91895 | -83.6435
0.91928 | -83.6488
0.91961 | -83.6446
0.91994 | -83.6355
0.92027 | -83.624
0.9206 | -83.6207
0.92093 | -83.6287
0.92126 | -83.6308
0.92159 | -83.6315
0.92192 | -83.6332
0.92224 | -83.636
0.92257 | -83.6394
0.9229 | -83.6433
0.92323 | -83.6467
0.92356 | -83.6499
0.92389 | -83.6547
0.92422 | -83.6636
0.92455 | -83.6722
0.92488 | -83.6819
0.92521 | -83.6895
0.92554 | -83.6929
0.92587 | -83.6962
0.92619 | -83.7013
0.92652 | -83.7063
0.92685 | -83.7082
0.92718 | -83.7158
0.92751 | -83.7242
0.92784 | -83.731
0.92817 | -83.734
0.9285 | -83.7369
0.92883 | -83.7422
0.92916 | -83.7505
0.92949 | -83.7623
0.92981 | -83.7714
0.93014 | -83.7809
0.93047 | -83.7883
0.9308 | -83.7952
0.93113 | -83.8018
0.93146 | -83.8115
0.93179 | -83.8195
0.93212 | -83.8192
0.93245 | -83.8219
0.93278 | -83.8273
0.93311 | -83.8281
0.93344 | -83.8319
0.93376 | -83.8363
0.93409 | -83.841
0.93442 | -83.8467
0.93475 | -83.8559
0.93508 | -83.8626
0.93541 | -83.8687
0.93574 | -83.8711
0.93607 | -83.871
0.9364 | -83.8791
0.93673 | -83.8874
0.93706 | -83.8968
0.93739 | -83.9019
0.93771 | -83.901
0.93804 | -83.9012
0.93837 | -83.905
0.9387 | -83.9082
0.93903 | -83.916
0.93936 | -83.9199
0.93969 | -83.919
0.94002 | -83.9228
0.94035 | -83.9263
0.94068 | -83.9311
0.941 | -83.9349
0.94133 | -83.9378
0.94166 | -83.9405
0.94199 | -83.9441
0.94232 | -83.9473
0.94265 | -83.9502
0.94298 | -83.9529
0.94331 | -83.9549
0.94364 | -83.9576
0.94397 | -83.9594
0.9443 | -83.964
0.94463 | -83.973
0.94495 | -83.9752
0.94528 | -83.9773
0.94561 | -83.9748
0.94594 | -83.9768
0.94627 | -83.9777
0.9466 | -83.9757
0.94693 | -83.9743
0.94726 | -83.9759
0.94759 | -83.9706
0.94792 | -83.9677
0.94825 | -83.9684
0.94857 | -83.9709
0.9489 | -83.9772
0.94923 | -83.9884
0.94956 | -83.9823
0.94989 | -83.9795
0.95022 | -83.9707
0.95055 | -83.9666
0.95088 | -83.9638
0.95121 | -83.9582
0.95154 | -83.9524
0.95187 | -83.9562
0.9522 | -83.9546
0.95252 | -83.9521
0.95285 | -83.9518
0.95318 | -83.9512
0.95351 | -83.9516
0.95384 | -83.9466
0.95417 | -83.9446
0.9545 | -83.9405
0.95483 | -83.9375
0.95516 | -83.9337
0.95549 | -83.9318
0.95582 | -83.9299
0.95615 | -83.9246
0.95647 | -83.9103
0.9568 | -83.9021
0.95713 | -83.8998
0.95746 | -83.8896
0.95779 | -83.8787
0.95812 | -83.8748
0.95845 | -83.8693
0.95878 | -83.8632
0.95911 | -83.8557
0.95944 | -83.8514
0.95977 | -83.8433
0.9601 | -83.8309
0.96042 | -83.8278
0.96075 | -83.8267
0.96108 | -83.8217
0.96141 | -83.8162
0.96174 | -83.8084
0.96207 | -83.8061
0.9624 | -83.8081
0.96273 | -83.8046
0.96306 | -83.7959
0.96339 | -83.7915
0.96372 | -83.7897
0.96405 | -83.7834
0.96437 | -83.7752
0.9647 | -83.7689
0.96503 | -83.7594
0.96536 | -83.7542
0.96569 | -83.7471
0.96602 | -83.7422
0.96635 | -83.709
0.96668 | -83.6841
0.96701 | -83.6751
0.96734 | -83.6675
0.96766 | -83.6593
0.968 | -83.6497
0.96832 | -83.6417
0.96865 | -83.6304
0.96898 | -83.6198
0.96931 | -83.6209
0.96964 | -83.622
0.96997 | -83.6149
0.9703 | -83.6038
0.97063 | -83.5943
0.97096 | -83.5805
0.97129 | -83.5713
0.97162 | -83.5591
0.97194 | -83.548
0.97227 | -83.538
0.9726 | -83.5268
0.97293 | -83.5149
0.97326 | -83.5029
0.97359 | -83.489
0.97392 | -83.4719
0.97425 | -83.4581
0.97458 | -83.4499
0.97491 | -83.4417
0.97524 | -83.4315
0.97557 | -83.4218
0.97589 | -83.4106
0.97622 | -83.3996
0.97655 | -83.3896
0.97688 | -83.3793
0.97721 | -83.3689
0.97754 | -83.3571
0.97787 | -83.339
0.9782 | -83.3232
0.97853 | -83.3131
0.97886 | -83.2998
0.97918 | -83.2914
0.97951 | -83.2815
0.97984 | -83.271
0.98017 | -83.2612
0.9805 | -83.2511
0.98083 | -83.2401
0.98116 | -83.2308
0.98149 | -83.2231
0.98182 | -83.2164
0.98215 | -83.2061
0.98248 | -83.1956
0.98281 | -83.1839
0.98313 | -83.1747
0.98346 | -83.1706
0.98379 | -83.1635
0.98412 | -83.1476
0.98445 | -83.1348
0.98478 | -83.1273
0.98511 | -83.1199
0.98544 | -83.1136
0.98577 | -83.1063
0.9861 | -83.1054
0.98643 | -83.1095
0.98676 | -83.1114
0.98708 | -83.1071
0.98741 | -83.1004
0.98774 | -83.0893
0.98807 | -83.0811
0.9884 | -83.0739
0.98873 | -83.069
0.98906 | -83.064
0.98939 | -83.0586
0.98972 | -83.0529
0.99005 | -83.0463
0.99038 | -83.042
0.9907 | -83.0382
0.99103 | -83.0355
0.99136 | -83.0327
0.99169 | -83.0296
0.99202 | -83.0278
0.99235 | -83.0237
0.99268 | -83.0192
0.99301 | -83.0161
0.99334 | -83.015
0.99367 | -83.0144
0.994 | -83.0166
0.99433 | -83.0168
0.99465 | -83.0163
0.99498 | -83.0166
0.99531 | -83.0162
0.99564 | -83.018
0.99597 | -83.0206
0.9963 | -83.0222
0.99663 | -83.022
0.99696 | -83.0279
0.99729 | -83.0236
0.99762 | -83.0254
0.99795 | -83.0288
0.99828 | -83.0342
0.9986 | -83.0375
0.99893 | -83.0415
0.99926 | -83.0469
0.99959 | -83.0527
0.99992 | -83.0617
1.0003 | -83.0671
1.0006 | -83.0725
1.0009 | -83.08
1.0012 | -83.086
1.0016 | -83.091
1.0019 | -83.0965
1.0022 | -83.1036
1.0026 | -83.1109
1.0029 | -83.1178
1.0032 | -83.1264
1.0035 | -83.1374
1.0039 | -83.15
1.0042 | -83.1596
1.0045 | -83.1674
1.0049 | -83.1732
1.0052 | -83.1776
1.0055 | -83.188
1.0058 | -83.1993
1.0062 | -83.2087
1.0065 | -83.2195
1.0068 | -83.225
1.0072 | -83.2359
1.0075 | -83.2468
1.0078 | -83.2593
1.0081 | -83.2696
1.0085 | -83.279
1.0088 | -83.2875
1.0091 | -83.2958
1.0095 | -83.3057
1.0098 | -83.3164
1.0101 | -83.3281
1.0105 | -83.339
1.0108 | -83.3488
1.0111 | -83.3598
1.0114 | -83.3724
1.0118 | -83.3861
1.0121 | -83.3999
1.0124 | -83.4142
1.0128 | -83.4302
1.0131 | -83.4417
1.0134 | -83.4562
1.0137 | -83.4701
1.0141 | -83.4888
1.0144 | -83.5063
1.0147 | -83.52
1.0151 | -83.5339
1.0154 | -83.5527
1.0157 | -83.5692
1.016 | -83.5823
1.0164 | -83.5956
1.0167 | -83.6072
1.017 | -83.6197
1.0174 | -83.6309
1.0177 | -83.6471
1.018 | -83.664
1.0184 | -83.68
1.0187 | -83.6981
1.019 | -83.7181
1.0193 | -83.7378
1.0197 | -83.7541
1.02 | -83.7697
1.0203 | -83.7891
1.0207 | -83.8057
1.021 | -83.8199
1.0213 | -83.8344
1.0216 | -83.8481
1.022 | -83.8637
1.0223 | -83.8777
1.0226 | -83.8925
1.023 | -83.9074
1.0233 | -83.9222
1.0236 | -83.9378
1.0239 | -83.9526
1.0243 | -83.968
1.0246 | -83.9813
1.0249 | -83.9954
1.0253 | -84.014
1.0256 | -84.0323
1.0259 | -84.0575
1.0263 | -84.0774
1.0266 | -84.0945
1.0269 | -84.1151
1.0272 | -84.1376
1.0276 | -84.1526
1.0279 | -84.1686
1.0282 | -84.1854
1.0286 | -84.202
1.0289 | -84.2166
1.0292 | -84.2299
1.0295 | -84.2503
1.0299 | -84.2738
1.0302 | -84.2898
1.0305 | -84.3021
1.0309 | -84.3096
1.0312 | -84.3128
1.0315 | -84.3192
1.0318 | -84.3288
1.0322 | -84.3424
1.0325 | -84.3571
1.0328 | -84.3721
1.0332 | -84.3859
1.0335 | -84.3998
1.0338 | -84.4121
1.0342 | -84.4231
1.0345 | -84.4344
1.0348 | -84.4444
1.0351 | -84.4581
1.0355 | -84.4665
1.0358 | -84.4736
1.0361 | -84.4822
1.0365 | -84.5233
1.0368 | -84.5496
1.0371 | -84.5547
1.0374 | -84.5605
1.0378 | -84.5636
1.0381 | -84.5693
1.0384 | -84.5749
1.0388 | -84.5818
1.0391 | -84.5911
1.0394 | -84.6013
1.0397 | -84.608
1.0401 | -84.6083
1.0404 | -84.608
1.0407 | -84.6147
1.0411 | -84.622
1.0414 | -84.6336
1.0417 | -84.6366
1.0421 | -84.6404
1.0424 | -84.6371
1.0427 | -84.6348
1.043 | -84.6383
1.0434 | -84.6429
1.0437 | -84.6463
1.044 | -84.654
1.0444 | -84.6492
1.0447 | -84.6533
1.045 | -84.6522
1.0453 | -84.6523
1.0457 | -84.6513
1.046 | -84.6506
1.0463 | -84.6495
1.0467 | -84.6494
1.047 | -84.6488
1.0473 | -84.6469
1.0476 | -84.6446
1.048 | -84.6404
1.0483 | -84.6376
1.0486 | -84.6358
1.049 | -84.6382
1.0493 | -84.6415
1.0496 | -84.6433
1.0499 | -84.6417
1.0503 | -84.6369
1.0506 | -84.6294
1.0509 | -84.622
1.0513 | -84.6151
1.0516 | -84.6082
1.0519 | -84.6043
1.0523 | -84.5954
1.0526 | -84.587
1.0529 | -84.5793
1.0532 | -84.5722
1.0536 | -84.566
1.0539 | -84.5568
1.0542 | -84.5518
1.0546 | -84.5459
1.0549 | -84.537
1.0552 | -84.5314
1.0555 | -84.5331
1.0559 | -84.5298
1.0562 | -84.5209
1.0565 | -84.5127
1.0569 | -84.5036
1.0572 | -84.4936
1.0575 | -84.4834
1.0578 | -84.4753
1.0582 | -84.4678
1.0585 | -84.461
1.0588 | -84.4512
1.0592 | -84.451
1.0595 | -84.4475
1.0598 | -84.4386
1.0602 | -84.428
1.0605 | -84.4162
1.0608 | -84.4083
1.0611 | -84.4
1.0615 | -84.3927
1.0618 | -84.3879
1.0621 | -84.3808
1.0625 | -84.3737
1.0628 | -84.3658
1.0631 | -84.3612
1.0634 | -84.3652
1.0638 | -84.3592
1.0641 | -84.3443
1.0644 | -84.3396
1.0648 | -84.3388
1.0651 | -84.336
1.0654 | -84.3336
1.0657 | -84.3276
1.0661 | -84.3213
1.0664 | -84.3187
1.0667 | -84.3168
1.0671 | -84.3153
1.0674 | -84.3137
1.0677 | -84.3073
1.0681 | -84.3044
1.0684 | -84.3034
1.0687 | -84.3052
1.069 | -84.3057
1.0694 | -84.3065
1.0697 | -84.3123
1.07 | -84.3095
1.0704 | -84.3098
1.0707 | -84.3089
1.071 | -84.3123
1.0713 | -84.3151
1.0717 | -84.3173
1.072 | -84.3213
1.0723 | -84.3251
1.0727 | -84.3302
1.073 | -84.3442
1.0733 | -84.3482
1.0736 | -84.3534
1.074 | -84.3584
1.0743 | -84.3613
1.0746 | -84.3667
1.075 | -84.3724
1.0753 | -84.377
1.0756 | -84.3834
1.076 | -84.3896
1.0763 | -84.3966
1.0766 | -84.4036
1.0769 | -84.4106
1.0773 | -84.4203
1.0776 | -84.4293
1.0779 | -84.4399
1.0783 | -84.4484
1.0786 | -84.4566
1.0789 | -84.4653
1.0792 | -84.4746
1.0796 | -84.4904
1.0799 | -84.5046
1.0802 | -84.5136
1.0806 | -84.5241
1.0809 | -84.5336
1.0812 | -84.5407
1.0815 | -84.5499
1.0819 | -84.5629
1.0822 | -84.5761
1.0825 | -84.5906
1.0829 | -84.6028
1.0832 | -84.6132
1.0835 | -84.6214
1.0839 | -84.6315
1.0842 | -84.6446
1.0845 | -84.6572
1.0848 | -84.668
1.0852 | -84.6803
1.0855 | -84.6917
1.0858 | -84.705
1.0862 | -84.7178
1.0865 | -84.7294
1.0868 | -84.7406
1.0871 | -84.7516
1.0875 | -84.762
1.0878 | -84.7742
1.0881 | -84.787
1.0885 | -84.8025
1.0888 | -84.8122
1.0891 | -84.822
1.0894 | -84.8353
1.0898 | -84.8495
1.0901 | -84.8664
1.0904 | -84.8816
1.0908 | -84.8933
1.0911 | -84.9076
1.0914 | -84.9184
1.0917 | -84.9255
1.0921 | -84.9308
1.0924 | -84.9399
1.0927 | -84.9503
1.0931 | -84.9601
1.0934 | -84.9653
1.0937 | -84.9673
1.0941 | -84.9696
1.0944 | -84.9714
1.0947 | -84.9734
1.095 | -84.981
1.0954 | -84.9886
1.0957 | -84.9967
1.096 | -84.9961
1.0964 | -85.0008
1.0967 | -85.0081
1.097 | -85.0159
1.0973 | -85.0222
1.0977 | -85.0278
1.098 | -85.0348
1.0983 | -85.0394
1.0987 | -85.0472
1.099 | -85.0559
1.0993 | -85.0644
1.0997 | -85.069
1.1 | -85.0731
1.1003 | -85.0722
1.1006 | -85.0684
1.101 | -85.0689
1.1013 | -85.0717
1.1016 | -85.0758
1.102 | -85.0788
1.1023 | -85.0864
1.1026 | -85.0888
1.1029 | -85.0918
1.1033 | -85.0957
1.1036 | -85.0886
1.1039 | -85.0814
1.1043 | -85.0773
1.1046 | -85.0752
1.1049 | -85.0752
1.1052 | -85.0786
1.1056 | -85.0872
1.1059 | -85.095
1.1062 | -85.0979
1.1066 | -85.0991
1.1069 | -85.0973
1.1072 | -85.0959
1.1075 | -85.0962
1.1079 | -85.0952
1.1082 | -85.0927
1.1085 | -85.094
1.1089 | -85.0952
1.1092 | -85.0914
1.1095 | -85.0877
1.1099 | -85.0871
1.1102 | -85.0835
1.1105 | -85.0837
1.1108 | -85.0743
1.1112 | -85.0729
1.1115 | -85.0683
1.1118 | -85.0611
1.1122 | -85.0551
1.1125 | -85.0512
1.1128 | -85.0475
1.1131 | -85.0439
1.1135 | -85.0331
1.1138 | -85.0198
1.1141 | -85.0079
1.1145 | -84.9975
1.1148 | -84.9867
1.1151 | -84.9806
1.1154 | -84.9778
1.1158 | -84.975
1.1161 | -84.9655
1.1164 | -84.9562
1.1168 | -84.9474
1.1171 | -84.9419
1.1174 | -84.936
1.1178 | -84.9334
1.1181 | -84.9317
1.1184 | -84.9226
1.1187 | -84.9092
1.1191 | -84.8992
1.1194 | -84.8879
1.1197 | -84.8749
1.1201 | -84.8528
1.1204 | -84.8385
1.1207 | -84.8321
1.121 | -84.8264
1.1214 | -84.8146
1.1217 | -84.804
1.122 | -84.793
1.1224 | -84.7824
1.1227 | -84.7715
1.123 | -84.7613
1.1233 | -84.7481
1.1237 | -84.7392
1.124 | -84.7264
1.1243 | -84.7174
1.1247 | -84.7101
1.125 | -84.7013
1.1253 | -84.6906
1.1257 | -84.6774
1.126 | -84.6652
1.1263 | -84.6546
1.1266 | -84.6458
1.127 | -84.6343
1.1273 | -84.6227
1.1276 | -84.6121
1.128 | -84.602
1.1283 | -84.5944
1.1286 | -84.5842
1.1289 | -84.5744
1.1293 | -84.5631
1.1296 | -84.5518
1.1299 | -84.5445
1.1303 | -84.5334
1.1306 | -84.5244
1.1309 | -84.5114
1.1312 | -84.4946
1.1316 | -84.4818
1.1319 | -84.4709
1.1322 | -84.4599
1.1326 | -84.4498
1.1329 | -84.438
1.1332 | -84.4233
1.1336 | -84.408
1.1339 | -84.3966
1.1342 | -84.3893
1.1345 | -84.3809
1.1349 | -84.3739
1.1352 | -84.3645
1.1355 | -84.3578
1.1359 | -84.3485
1.1362 | -84.3376
1.1365 | -84.3279
1.1368 | -84.3184
1.1372 | -84.3104
1.1375 | -84.3034
1.1378 | -84.2948
1.1382 | -84.2885
1.1385 | -84.2821
1.1388 | -84.274
1.1391 | -84.2636
1.1395 | -84.2502
1.1398 | -84.2378
1.1401 | -84.2249
1.1405 | -84.2106
1.1408 | -84.1994
1.1411 | -84.1906
1.1415 | -84.1923
1.1418 | -84.1875
1.1421 | -84.182
1.1424 | -84.1765
1.1428 | -84.1695
1.1431 | -84.1663
1.1434 | -84.1627
1.1438 | -84.157
1.1441 | -84.1512
1.1444 | -84.1455
1.1447 | -84.1437
1.1451 | -84.144
1.1454 | -84.1291
1.1457 | -84.119
1.1461 | -84.1141
1.1464 | -84.1103
1.1467 | -84.1083
1.147 | -84.1049
1.1474 | -84.1041
1.1477 | -84.1001
1.148 | -84.0968
1.1484 | -84.0937
1.1487 | -84.083
1.149 | -84.0672
1.1494 | -84.0589
1.1497 | -84.0522
1.15 | -84.0464
1.1503 | -84.038
1.1507 | -84.0285
1.151 | -84.0199
1.1513 | -84.0134
1.1517 | -84.0057
1.152 | -84.0014
1.1523 | -83.9988
1.1526 | -83.9932
1.153 | -83.9875
1.1533 | -83.9825
1.1536 | -83.977
1.154 | -83.9722
1.1543 | -83.9669
1.1546 | -83.963
1.1549 | -83.9595
1.1553 | -83.9534
1.1556 | -83.9495
1.1559 | -83.9487
1.1563 | -83.9458
1.1566 | -83.9406
1.1569 | -83.9336
1.1572 | -83.9246
1.1576 | -83.9155
1.1579 | -83.9102
1.1582 | -83.9049
1.1586 | -83.9003
1.1589 | -83.8958
1.1592 | -83.8887
1.1596 | -83.8849
1.1599 | -83.8816
1.1602 | -83.8768
1.1605 | -83.8708
1.1609 | -83.8634
1.1612 | -83.8577
1.1615 | -83.848
1.1619 | -83.8419
1.1622 | -83.839
1.1625 | -83.8324
1.1628 | -83.8263
1.1632 | -83.8165
1.1635 | -83.8105
1.1638 | -83.8046
1.1642 | -83.7999
1.1645 | -83.7938
1.1648 | -83.7835
1.1651 | -83.7731
1.1655 | -83.7671
1.1658 | -83.7612
1.1661 | -83.7557
1.1665 | -83.7498
1.1668 | -83.7412
1.1671 | -83.7298
1.1675 | -83.7179
1.1678 | -83.7074
1.1681 | -83.7024
1.1684 | -83.6961
1.1688 | -83.6904
1.1691 | -83.6843
1.1694 | -83.679
1.1698 | -83.6733
1.1701 | -83.6676
1.1704 | -83.6611
1.1707 | -83.6495
1.1711 | -83.6403
1.1714 | -83.6341
1.1717 | -83.6287
1.1721 | -83.6226
1.1724 | -83.6155
1.1727 | -83.6099
1.173 | -83.6069
1.1734 | -83.6023
1.1737 | -83.6
1.174 | -83.5972
1.1744 | -83.5883
1.1747 | -83.582
1.175 | -83.5768
1.1754 | -83.5759
1.1757 | -83.5726
1.176 | -83.567
1.1763 | -83.5617
1.1767 | -83.5551
1.177 | -83.5485
1.1773 | -83.5473
1.1777 | -83.543
1.178 | -83.5397
1.1783 | -83.5355
1.1786 | -83.5308
1.179 | -83.5262
1.1793 | -83.5216
1.1796 | -83.5188
1.18 | -83.5191
1.1803 | -83.5197
1.1806 | -83.5173
1.1809 | -83.5129
1.1813 | -83.5081
1.1816 | -83.5042
1.1819 | -83.4999
1.1823 | -83.4947
1.1826 | -83.4921
1.1829 | -83.4908
1.1833 | -83.4903
1.1836 | -83.4944
1.1839 | -83.4994
1.1842 | -83.5015
1.1846 | -83.5013
1.1849 | -83.5015
1.1852 | -83.5008
1.1856 | -83.4989
1.1859 | -83.5002
1.1862 | -83.5035
1.1865 | -83.5046
1.1869 | -83.5121
1.1872 | -83.5124
1.1875 | -83.513
1.1879 | -83.5147
1.1882 | -83.5159
1.1885 | -83.517
1.1888 | -83.5188
1.1892 | -83.5214
1.1895 | -83.5248
1.1898 | -83.5239
1.1902 | -83.5182
1.1905 | -83.5208
1.1908 | -83.5242
1.1911 | -83.5291
1.1915 | -83.5365
1.1918 | -83.5426
1.1921 | -83.5484
1.1925 | -83.5544
1.1928 | -83.5605
1.1931 | -83.5665
1.1935 | -83.5734
1.1938 | -83.5795
1.1941 | -83.589
1.1944 | -83.5993
1.1948 | -83.6055
1.1951 | -83.6124
1.1954 | -83.62
1.1958 | -83.6258
1.1961 | -83.6326
1.1964 | -83.645
1.1967 | -83.6529
1.1971 | -83.6603
1.1974 | -83.6663
1.1977 | -83.6711
1.1981 | -83.6782
1.1984 | -83.6871
1.1987 | -83.6959
1.199 | -83.7023
1.1994 | -83.7108
1.1997 | -83.7221
1.2 | -83.7322
1.2004 | -83.7421
1.2007 | -83.7546
1.201 | -83.7706
1.2014 | -83.7779
1.2017 | -83.7837
1.202 | -83.7928
1.2023 | -83.8017
1.2027 | -83.8113
1.203 | -83.8191
1.2033 | -83.8261
1.2037 | -83.832
1.204 | -83.8403
1.2043 | -83.8503
1.2046 | -83.8629
1.205 | -83.8732
1.2053 | -83.8869
1.2056 | -83.8963
1.206 | -83.91
1.2063 | -83.9214
1.2066 | -83.9312
1.2069 | -83.9399
1.2073 | -83.9494
1.2076 | -83.959
1.2079 | -83.9684
1.2083 | -83.9794
1.2086 | -83.9919
1.2089 | -84.0045
1.2093 | -84.0162
1.2096 | -84.0261
1.2099 | -84.0371
1.2102 | -84.0496
1.2106 | -84.0612
1.2109 | -84.0738
1.2112 | -84.0835
1.2116 | -84.0942
1.2119 | -84.1077
1.2122 | -84.1193
1.2125 | -84.1318
1.2129 | -84.1472
1.2132 | -84.165
1.2135 | -84.1795
1.2139 | -84.192
1.2142 | -84.2045
1.2145 | -84.2174
1.2148 | -84.2315
1.2152 | -84.2459
1.2155 | -84.2567
1.2158 | -84.2731
1.2162 | -84.2869
1.2165 | -84.2991
1.2168 | -84.3103
1.2172 | -84.3228
1.2175 | -84.3341
1.2178 | -84.3435
1.2181 | -84.355
1.2185 | -84.3692
1.2188 | -84.383
1.2191 | -84.3953
1.2195 | -84.405
1.2198 | -84.4138
1.2201 | -84.4256
1.2204 | -84.4338
1.2208 | -84.4428
1.2211 | -84.455
1.2214 | -84.4693
1.2218 | -84.4855
1.2221 | -84.5008
1.2224 | -84.5154
1.2227 | -84.531
1.2231 | -84.5479
1.2234 | -84.562
1.2237 | -84.5751
1.2241 | -84.5863
1.2244 | -84.596
1.2247 | -84.6064
1.2251 | -84.6148
1.2254 | -84.6226
1.2257 | -84.6287
1.226 | -84.6289
1.2264 | -84.6335
1.2267 | -84.6428
1.227 | -84.6511
1.2274 | -84.6593
1.2277 | -84.6659
1.228 | -84.6744
1.2283 | -84.6935
1.2287 | -84.7056
1.229 | -84.7119
1.2293 | -84.7186
1.2297 | -84.7267
1.23 | -84.7347
1.2303 | -84.7427
1.2306 | -84.7496
1.231 | -84.757
1.2313 | -84.765
1.2316 | -84.7798
1.232 | -84.7979
1.2323 | -84.803
1.2326 | -84.8103
1.233 | -84.818
1.2333 | -84.8257
1.2336 | -84.8367
1.2339 | -84.8472
1.2343 | -84.85
1.2346 | -84.8584
1.2349 | -84.8673
1.2353 | -84.8749
1.2356 | -84.8807
1.2359 | -84.8831
1.2362 | -84.8881
1.2366 | -84.8943
1.2369 | -84.8976
1.2372 | -84.8991
1.2376 | -84.9039
1.2379 | -84.9087
1.2382 | -84.9137
1.2385 | -84.9146
1.2389 | -84.919
1.2392 | -84.9227
1.2395 | -84.9265
1.2399 | -84.9294
1.2402 | -84.9257
1.2405 | -84.9244
1.2409 | -84.9323
1.2412 | -84.9394
1.2415 | -84.9432
1.2418 | -84.9464
1.2422 | -84.9497
1.2425 | -84.9527
1.2428 | -84.9525
1.2432 | -84.9559
1.2435 | -84.9593
1.2438 | -84.9584
1.2441 | -84.9554
1.2445 | -84.9515
1.2448 | -84.9529
1.2451 | -84.9562
1.2455 | -84.9599
1.2458 | -84.9659
1.2461 | -84.9653
1.2464 | -84.9607
1.2468 | -84.9629
1.2471 | -84.9648
1.2474 | -84.9682
1.2478 | -84.9664
1.2481 | -84.9642
1.2484 | -84.9622
1.2487 | -84.9597
1.2491 | -84.9651
1.2494 | -84.9719
1.2497 | -84.9718
1.2501 | -84.9693
1.2504 | -84.9657
1.2507 | -84.9614
1.2511 | -84.9633
1.2514 | -84.969
1.2517 | -84.9752
1.252 | -84.9767
1.2524 | -84.9744
1.2527 | -84.9725
1.253 | -84.9707
1.2534 | -84.969
1.2537 | -84.96
1.254 | -84.9486
1.2543 | -84.943
1.2547 | -84.9379
1.255 | -84.932
1.2553 | -84.93
1.2557 | -84.9369
1.256 | -84.9412
1.2563 | -84.9385
1.2566 | -84.9352
1.257 | -84.9318
1.2573 | -84.9302
1.2576 | -84.929
1.258 | -84.928
1.2583 | -84.9282
1.2586 | -84.9261
1.259 | -84.9252
1.2593 | -84.9253
1.2596 | -84.9235
1.2599 | -84.9217
1.2603 | -84.9183
1.2606 | -84.9138
1.2609 | -84.9099
1.2613 | -84.9065
1.2616 | -84.9048
1.2619 | -84.9028
1.2622 | -84.8966
1.2626 | -84.8905
1.2629 | -84.8835
1.2632 | -84.8797
1.2636 | -84.8767
1.2639 | -84.8717
1.2642 | -84.8644
1.2645 | -84.8599
1.2649 | -84.8562
1.2652 | -84.8497
1.2655 | -84.8471
1.2659 | -84.8458
1.2662 | -84.8431
1.2665 | -84.8374
1.2669 | -84.8349
1.2672 | -84.8297
1.2675 | -84.8258
1.2678 | -84.8247
1.2682 | -84.821
1.2685 | -84.814
1.2688 | -84.8095
1.2692 | -84.8068
1.2695 | -84.8023
1.2698 | -84.7974
1.2701 | -84.7951
1.2705 | -84.7892
1.2708 | -84.78
1.2711 | -84.7746
1.2715 | -84.7683
1.2718 | -84.7607
1.2721 | -84.7547
1.2724 | -84.7489
1.2728 | -84.7434
1.2731 | -84.7384
1.2734 | -84.7339
1.2738 | -84.7295
1.2741 | -84.7235
1.2744 | -84.7188
1.2748 | -84.716
1.2751 | -84.711
1.2754 | -84.7072
1.2757 | -84.7026
1.2761 | -84.7034
1.2764 | -84.7092
1.2767 | -84.7054
1.2771 | -84.6995
1.2774 | -84.6951
1.2777 | -84.6895
1.278 | -84.6852
1.2784 | -84.6806
1.2787 | -84.6779
1.279 | -84.6743
1.2794 | -84.6713
1.2797 | -84.6669
1.28 | -84.6651
1.2803 | -84.662
1.2807 | -84.6576
1.281 | -84.6557
1.2813 | -84.6542
1.2817 | -84.651
1.282 | -84.6498
1.2823 | -84.6466
1.2827 | -84.6447
1.283 | -84.6411
1.2833 | -84.6376
1.2836 | -84.6332
1.284 | -84.6301
1.2843 | -84.6297
1.2846 | -84.6287
1.285 | -84.6271
1.2853 | -84.6247
1.2856 | -84.6223
1.2859 | -84.6216
1.2863 | -84.6201
1.2866 | -84.6173
1.2869 | -84.6123
1.2873 | -84.6114
1.2876 | -84.6111
1.2879 | -84.6101
1.2882 | -84.6106
1.2886 | -84.612
1.2889 | -84.6117
1.2892 | -84.6126
1.2896 | -84.6137
1.2899 | -84.6155
1.2902 | -84.6171
1.2906 | -84.6208
1.2909 | -84.6229
1.2912 | -84.6258
1.2915 | -84.6192
1.2919 | -84.6124
1.2922 | -84.6039
1.2925 | -84.6005
1.2929 | -84.6031
1.2932 | -84.603
1.2935 | -84.6004
1.2938 | -84.5982
1.2942 | -84.5979
1.2945 | -84.5999
1.2948 | -84.6036
1.2952 | -84.6064
1.2955 | -84.6082
1.2958 | -84.6093
1.2961 | -84.6106
1.2965 | -84.6093
1.2968 | -84.6055
1.2971 | -84.5994
1.2975 | -84.5939
1.2978 | -84.5917
1.2981 | -84.5935
1.2984 | -84.5883
1.2988 | -84.5869
1.2991 | -84.5888
1.2994 | -84.5882
1.2998 | -84.5863
1.3001 | -84.5846
1.3004 | -84.5831
1.3008 | -84.5838
1.3011 | -84.5805
1.3014 | -84.5767
1.3017 | -84.5758
1.3021 | -84.5747
1.3024 | -84.574
1.3027 | -84.5717
1.3031 | -84.5692
1.3034 | -84.563
1.3037 | -84.5624
1.304 | -84.563
1.3044 | -84.5601
1.3047 | -84.5542
1.305 | -84.5498
1.3054 | -84.5467
1.3057 | -84.5444
1.306 | -84.5442
1.3063 | -84.5432
1.3067 | -84.5393
1.307 | -84.5364
1.3073 | -84.5376
1.3077 | -84.5395
1.308 | -84.5388
1.3083 | -84.5366
1.3087 | -84.5368
1.309 | -84.534
1.3093 | -84.5299
1.3096 | -84.5279
1.31 | -84.5251
1.3103 | -84.5131
1.3106 | -84.5124
1.311 | -84.5102
1.3113 | -84.5039
1.3116 | -84.4979
1.3119 | -84.4937
1.3123 | -84.4899
1.3126 | -84.4846
1.3129 | -84.478
1.3133 | -84.4711
1.3136 | -84.4666
1.3139 | -84.4623
1.3142 | -84.4584
1.3146 | -84.4521
1.3149 | -84.4477
1.3152 | -84.4427
1.3156 | -84.4364
1.3159 | -84.4304
1.3162 | -84.4246
1.3166 | -84.4208
1.3169 | -84.4145
1.3172 | -84.4018
1.3175 | -84.3884
1.3179 | -84.3793
1.3182 | -84.3718
1.3185 | -84.364
1.3189 | -84.3576
1.3192 | -84.3497
1.3195 | -84.3424
1.3198 | -84.333
1.3202 | -84.3271
1.3205 | -84.319
1.3208 | -84.3078
1.3212 | -84.297
1.3215 | -84.2926
1.3218 | -84.2859
1.3221 | -84.2771
1.3225 | -84.2668
1.3228 | -84.2543
1.3231 | -84.2449
1.3235 | -84.2371
1.3238 | -84.2299
1.3241 | -84.2218
1.3245 | -84.2127
1.3248 | -84.204
1.3251 | -84.1937
1.3254 | -84.1855
1.3258 | -84.1769
1.3261 | -84.1679
1.3264 | -84.1609
1.3268 | -84.1524
1.3271 | -84.1409
1.3274 | -84.1297
1.3277 | -84.1171
1.3281 | -84.11
1.3284 | -84.1019
1.3287 | -84.0933
1.3291 | -84.0832
1.3294 | -84.0704
1.3297 | -84.0601
1.33 | -84.0571
1.3304 | -84.0538
1.3307 | -84.0443
1.331 | -84.0352
1.3314 | -84.0263
1.3317 | -84.014
1.332 | -84.0024
1.3323 | -83.9934
1.3327 | -83.9864
1.333 | -83.9789
1.3333 | -83.9705
1.3337 | -83.9615
1.334 | -83.9546
1.3343 | -83.9496
1.3347 | -83.9446
1.335 | -83.9376
1.3353 | -83.9297
1.3356 | -83.9213
1.336 | -83.9146
1.3363 | -83.9105
1.3366 | -83.9047
1.337 | -83.8962
1.3373 | -83.8884
1.3376 | -83.8817
1.3379 | -83.8763
1.3383 | -83.8709
1.3386 | -83.865
1.3389 | -83.8569
1.3393 | -83.8519
1.3396 | -83.8464
1.3399 | -83.8393
1.3402 | -83.8332
1.3406 | -83.8295
1.3409 | -83.827
1.3412 | -83.8248
1.3416 | -83.8203
1.3419 | -83.8162
1.3422 | -83.8127
1.3426 | -83.8083
1.3429 | -83.8047
1.3432 | -83.8013
1.3435 | -83.799
1.3439 | -83.7942
1.3442 | -83.7884
1.3445 | -83.7895
1.3449 | -83.7939
1.3452 | -83.7916
1.3455 | -83.7858
1.3458 | -83.7827
1.3462 | -83.7797
1.3465 | -83.7778
1.3468 | -83.7758
1.3472 | -83.7748
1.3475 | -83.7742
1.3478 | -83.7727
1.3481 | -83.7673
1.3485 | -83.7737
1.3488 | -83.7746
1.3491 | -83.7755
1.3495 | -83.7749
1.3498 | -83.7746
1.3501 | -83.7756
1.3505 | -83.7766
1.3508 | -83.7765
1.3511 | -83.7775
1.3514 | -83.7798
1.3518 | -83.7823
1.3521 | -83.7832
1.3524 | -83.785
1.3528 | -83.784
1.3531 | -83.784
1.3534 | -83.791
1.3537 | -83.7971
1.3541 | -83.803
1.3544 | -83.8085
1.3547 | -83.813
1.3551 | -83.8168
1.3554 | -83.8212
1.3557 | -83.8248
1.356 | -83.8299
1.3564 | -83.8357
1.3567 | -83.8424
1.357 | -83.8475
1.3574 | -83.8488
1.3577 | -83.8581
1.358 | -83.8626
1.3584 | -83.8698
1.3587 | -83.8765
1.359 | -83.8853
1.3593 | -83.8926
1.3597 | -83.8988
1.36 | -83.9059
1.3603 | -83.9144
1.3607 | -83.9212
1.361 | -83.9283
1.3613 | -83.9326
1.3616 | -83.9404
1.362 | -83.9476
1.3623 | -83.9567
1.3626 | -83.9684
1.363 | -83.9783
1.3633 | -83.9894
1.3636 | -84.0005
1.3639 | -84.0081
1.3643 | -84.0173
1.3646 | -84.0244
1.3649 | -84.0312
1.3653 | -84.04
1.3656 | -84.0463
1.3659 | -84.0529
1.3663 | -84.0618
1.3666 | -84.0683
1.3669 | -84.0731
1.3672 | -84.0825
1.3676 | -84.0911
1.3679 | -84.0995
1.3682 | -84.1071
1.3686 | -84.1115
1.3689 | -84.1188
1.3692 | -84.1279
1.3695 | -84.1374
1.3699 | -84.1438
1.3702 | -84.1509
1.3705 | -84.1581
1.3709 | -84.1656
1.3712 | -84.171
1.3715 | -84.1772
1.3718 | -84.1846
1.3722 | -84.1933
1.3725 | -84.2012
1.3728 | -84.2072
1.3732 | -84.2165
1.3735 | -84.2247
1.3738 | -84.2299
1.3742 | -84.2381
1.3745 | -84.2471
1.3748 | -84.2527
1.3751 | -84.2589
1.3755 | -84.2661
1.3758 | -84.275
1.3761 | -84.2834
1.3765 | -84.2903
1.3768 | -84.297
1.3771 | -84.3058
1.3774 | -84.3132
1.3778 | -84.3223
1.3781 | -84.329
1.3784 | -84.3364
1.3788 | -84.344
1.3791 | -84.3515
1.3794 | -84.3584
1.3797 | -84.3662
1.3801 | -84.3758
1.3804 | -84.3817
1.3807 | -84.3868
1.3811 | -84.395
1.3814 | -84.4052
1.3817 | -84.4098
1.3821 | -84.4068
1.3824 | -84.4058
1.3827 | -84.4106
1.383 | -84.4166
1.3834 | -84.4235
1.3837 | -84.4294
1.384 | -84.435
1.3844 | -84.4404
1.3847 | -84.4487
1.385 | -84.46
1.3853 | -84.4714
1.3857 | -84.4841
1.386 | -84.4909
1.3863 | -84.4951
1.3867 | -84.4977
1.387 | -84.5026
1.3873 | -84.5092
1.3876 | -84.5157
1.388 | -84.5218
1.3883 | -84.527
1.3886 | -84.5316
1.389 | -84.5341
1.3893 | -84.5381
1.3896 | -84.5425
1.3899 | -84.5462
1.3903 | -84.5514
1.3906 | -84.5562
1.3909 | -84.5681
1.3913 | -84.5786
1.3916 | -84.5845
1.3919 | -84.5836
1.3923 | -84.5912
1.3926 | -84.5995
1.3929 | -84.6008
1.3932 | -84.605
1.3936 | -84.6064
1.3939 | -84.6104
1.3942 | -84.6143
1.3946 | -84.6111
1.3949 | -84.6083
1.3952 | -84.6142
1.3955 | -84.6214
1.3959 | -84.6264
1.3962 | -84.6302
1.3965 | -84.6314
1.3969 | -84.6342
1.3972 | -84.6395
1.3975 | -84.6465
1.3978 | -84.6536
1.3982 | -84.657
1.3985 | -84.6597
1.3988 | -84.6649
1.3992 | -84.672
1.3995 | -84.6761
1.3998 | -84.6785
1.4002 | -84.683
1.4005 | -84.6852
1.4008 | -84.688
1.4011 | -84.69
1.4015 | -84.6903
1.4018 | -84.6906
1.4021 | -84.6928
1.4025 | -84.6897
1.4028 | -84.6866
1.4031 | -84.6878
1.4034 | -84.6898
1.4038 | -84.69
1.4041 | -84.6913
1.4044 | -84.6924
1.4048 | -84.6962
1.4051 | -84.6986
1.4054 | -84.7014
1.4057 | -84.7019
1.4061 | -84.703
1.4064 | -84.7045
1.4067 | -84.7078
1.4071 | -84.7096
1.4074 | -84.7099
1.4077 | -84.7122
1.4081 | -84.7168
1.4084 | -84.7215
1.4087 | -84.7234
1.409 | -84.7262
1.4094 | -84.7282
1.4097 | -84.7279
1.41 | -84.7307
1.4104 | -84.7349
1.4107 | -84.7347
1.411 | -84.7359
1.4113 | -84.7357
1.4117 | -84.7353
1.412 | -84.7371
1.4123 | -84.7414
1.4127 | -84.7437
1.413 | -84.7381
1.4133 | -84.7374
1.4136 | -84.7377
1.414 | -84.7366
1.4143 | -84.7359
1.4146 | -84.7347
1.415 | -84.728
1.4153 | -84.7265
1.4156 | -84.7269
1.416 | -84.725
1.4163 | -84.7228
1.4166 | -84.7254
1.4169 | -84.728
1.4173 | -84.7284
1.4176 | -84.7284
1.4179 | -84.7279
1.4183 | -84.7284
1.4186 | -84.7293
1.4189 | -84.7273
1.4192 | -84.7247
1.4196 | -84.7235
1.4199 | -84.721
1.4202 | -84.7146
1.4206 | -84.711
1.4209 | -84.7106
1.4212 | -84.7105
1.4215 | -84.7052
1.4219 | -84.6975
1.4222 | -84.6937
1.4225 | -84.6934
1.4229 | -84.6925
1.4232 | -84.6922
1.4235 | -84.6912
1.4239 | -84.6882
1.4242 | -84.6846
1.4245 | -84.6813
1.4248 | -84.6754
1.4252 | -84.6664
1.4255 | -84.6613
1.4258 | -84.6604
1.4262 | -84.6579
1.4265 | -84.6526
1.4268 | -84.6475
1.4271 | -84.6443
1.4275 | -84.6398
1.4278 | -84.6353
1.4281 | -84.6342
1.4285 | -84.6318
1.4288 | -84.6272
1.4291 | -84.6226
1.4294 | -84.6185
1.4298 | -84.612
1.4301 | -84.609
1.4304 | -84.605
1.4308 | -84.6004
1.4311 | -84.5961
1.4314 | -84.5912
1.4318 | -84.5849
1.4321 | -84.5796
1.4324 | -84.5738
1.4327 | -84.5661
1.4331 | -84.5634
1.4334 | -84.5607
1.4337 | -84.554
1.4341 | -84.5498
1.4344 | -84.5447
1.4347 | -84.5404
1.435 | -84.5334
1.4354 | -84.5258
1.4357 | -84.524
1.436 | -84.5212
1.4364 | -84.516
1.4367 | -84.5103
1.437 | -84.5064
1.4373 | -84.5014
1.4377 | -84.4947
1.438 | -84.488
1.4383 | -84.4791
1.4387 | -84.473
1.439 | -84.4656
1.4393 | -84.4577
1.4396 | -84.4542
1.44 | -84.4505
1.4403 | -84.4454
1.4406 | -84.4397
1.441 | -84.4359
1.4413 | -84.4347
1.4416 | -84.4315
1.442 | -84.4267
1.4423 | -84.4236
1.4426 | -84.4228
1.4429 | -84.4216
1.4433 | -84.4199
1.4436 | -84.4149
1.4439 | -84.4106
1.4443 | -84.4047
1.4446 | -84.396
1.4449 | -84.3891
1.4452 | -84.3868
1.4456 | -84.388
1.4459 | -84.3821
1.4462 | -84.3757
1.4466 | -84.3702
1.4469 | -84.3645
1.4472 | -84.3623
1.4475 | -84.3591
1.4479 | -84.3557
1.4482 | -84.3535
1.4485 | -84.3509
1.4489 | -84.3441
1.4492 | -84.3354
1.4495 | -84.3309
1.4499 | -84.3306
1.4502 | -84.3291
1.4505 | -84.3296
1.4508 | -84.3273
1.4512 | -84.3205
1.4515 | -84.3169
1.4518 | -84.3146
1.4522 | -84.3116
1.4525 | -84.3091
1.4528 | -84.3123
1.4531 | -84.3089
1.4535 | -84.3094
1.4538 | -84.3119
1.4541 | -84.3093
1.4545 | -84.308
1.4548 | -84.308
1.4551 | -84.3075
1.4554 | -84.311
1.4558 | -84.3124
1.4561 | -84.3094
1.4564 | -84.3113
1.4568 | -84.311
1.4571 | -84.3102
1.4574 | -84.3095
1.4578 | -84.3107
1.4581 | -84.3117
1.4584 | -84.3115
1.4587 | -84.3127
1.4591 | -84.3117
1.4594 | -84.3118
1.4597 | -84.317
1.4601 | -84.3188
1.4604 | -84.3177
1.4607 | -84.3148
1.461 | -84.3161
1.4614 | -84.3186
1.4617 | -84.3208
1.462 | -84.3228
1.4624 | -84.3286
1.4627 | -84.3319
1.463 | -84.3354
1.4633 | -84.3386
1.4637 | -84.3391
1.464 | -84.3343
1.4643 | -84.3365
1.4647 | -84.34
1.465 | -84.3433
1.4653 | -84.347
1.4657 | -84.3527
1.466 | -84.3539
1.4663 | -84.3579
1.4666 | -84.3617
1.467 | -84.3634
1.4673 | -84.3687
1.4676 | -84.3723
1.468 | -84.3723
1.4683 | -84.3739
1.4686 | -84.3754
1.4689 | -84.3805
1.4693 | -84.3866
1.4696 | -84.389
1.4699 | -84.3919
1.4703 | -84.3969
1.4706 | -84.4015
1.4709 | -84.4063
1.4712 | -84.4093
1.4716 | -84.4125
1.4719 | -84.4166
1.4722 | -84.4207
1.4726 | -84.4244
1.4729 | -84.4297
1.4732 | -84.4346
1.4735 | -84.4348
1.4739 | -84.4351
1.4742 | -84.4395
1.4745 | -84.4449
1.4749 | -84.448
1.4752 | -84.4502
1.4755 | -84.4551
1.4759 | -84.4594
1.4762 | -84.4628
1.4765 | -84.4664
1.4768 | -84.4704
1.4772 | -84.4717
1.4775 | -84.4745
1.4778 | -84.4792
1.4782 | -84.4821
1.4785 | -84.4857
1.4788 | -84.4857
1.4791 | -84.4859
1.4795 | -84.483
1.4798 | -84.4776
1.4801 | -84.4809
1.4805 | -84.4848
1.4808 | -84.4874
1.4811 | -84.4834
1.4814 | -84.4853
1.4818 | -84.491
1.4821 | -84.5013
1.4824 | -84.5074
1.4828 | -84.5097
1.4831 | -84.5119
1.4834 | -84.5144
1.4838 | -84.5142
1.4841 | -84.5123
1.4844 | -84.5144
1.4847 | -84.5163
1.4851 | -84.5187
1.4854 | -84.5187
1.4857 | -84.5171
1.4861 | -84.5205
1.4864 | -84.5216
1.4867 | -84.5241
1.487 | -84.5268
1.4874 | -84.526
1.4877 | -84.5215
1.488 | -84.5184
1.4884 | -84.5161
1.4887 | -84.5139
1.489 | -84.5134
1.4893 | -84.5086
1.4897 | -84.5051
1.49 | -84.5036
1.4903 | -84.5035
1.4907 | -84.498
1.491 | -84.4978
1.4913 | -84.4975
1.4917 | -84.4947
1.492 | -84.4931
1.4923 | -84.4907
1.4926 | -84.4926
1.493 | -84.4935
1.4933 | -84.4905
1.4936 | -84.4887
1.494 | -84.4866
1.4943 | -84.4888
1.4946 | -84.4864
1.4949 | -84.4849
1.4953 | -84.4889
1.4956 | -84.4899
1.4959 | -84.483
1.4963 | -84.4782
1.4966 | -84.4751
1.4969 | -84.4704
1.4972 | -84.4653
1.4976 | -84.4609
1.4979 | -84.4553
1.4982 | -84.4496
1.4986 | -84.446
1.4989 | -84.4435
1.4992 | -84.4384
1.4996 | -84.4336
1.4999 | -84.4285
1.5002 | -84.4156
1.5005 | -84.411
1.5009 | -84.4084
1.5012 | -84.4047
1.5015 | -84.4005
1.5019 | -84.3963
1.5022 | -84.3891
1.5025 | -84.3861
1.5028 | -84.3804
1.5032 | -84.3745
1.5035 | -84.3694
1.5038 | -84.369
1.5042 | -84.3657
1.5045 | -84.3592
1.5048 | -84.3551
1.5051 | -84.3523
1.5055 | -84.3471
1.5058 | -84.3413
1.5061 | -84.3367
1.5065 | -84.3344
1.5068 | -84.3291
1.5071 | -84.3226
1.5075 | -84.3176
1.5078 | -84.3123
1.5081 | -84.3055
1.5084 | -84.2937
1.5088 | -84.2785
1.5091 | -84.2666
1.5094 | -84.2585
1.5098 | -84.2405
1.5101 | -84.2315
1.5104 | -84.2245
1.5107 | -84.2173
1.5111 | -84.2083
1.5114 | -84.1991
1.5117 | -84.1925
1.5121 | -84.1827
1.5124 | -84.1719
1.5127 | -84.162
1.513 | -84.1489
1.5134 | -84.1408
1.5137 | -84.1334
1.514 | -84.1225
1.5144 | -84.1161
1.5147 | -84.1108
1.515 | -84.1067
1.5154 | -84.0951
1.5157 | -84.085
1.516 | -84.0773
1.5163 | -84.0687
1.5167 | -84.0602
1.517 | -84.0539
1.5173 | -84.0488
1.5177 | -84.0408
1.518 | -84.0322
1.5183 | -84.0236
1.5186 | -84.0162
1.519 | -84.0033
1.5193 | -83.9959
1.5196 | -83.9885
1.52 | -83.9789
1.5203 | -83.9694
1.5206 | -83.961
1.5209 | -83.9523
1.5213 | -83.9453
1.5216 | -83.9359
1.5219 | -83.9283
1.5223 | -83.9208
1.5226 | -83.9137
1.5229 | -83.9106
1.5232 | -83.9052
1.5236 | -83.8981
1.5239 | -83.8889
1.5242 | -83.8813
1.5246 | -83.8773
1.5249 | -83.8711
1.5252 | -83.8618
1.5256 | -83.8525
1.5259 | -83.8438
1.5262 | -83.8384
1.5265 | -83.8323
1.5269 | -83.826
1.5272 | -83.8177
1.5275 | -83.8115
1.5279 | -83.8074
1.5282 | -83.8031
1.5285 | -83.7957
1.5288 | -83.7928
1.5292 | -83.7877
1.5295 | -83.7818
1.5298 | -83.7767
1.5302 | -83.7711
1.5305 | -83.7649
1.5308 | -83.7593
1.5312 | -83.75
1.5315 | -83.7391
1.5318 | -83.7327
1.5321 | -83.7268
1.5325 | -83.7185
1.5328 | -83.7096
1.5331 | -83.7017
1.5335 | -83.6912
1.5338 | -83.6832
1.5341 | -83.6774
1.5344 | -83.6699
1.5348 | -83.6625
1.5351 | -83.6586
1.5354 | -83.6537
1.5358 | -83.6477
1.5361 | -83.6437
1.5364 | -83.6385
1.5367 | -83.6336
1.5371 | -83.63
1.5374 | -83.6266
1.5377 | -83.6222
1.5381 | -83.6189
1.5384 | -83.6119
1.5387 | -83.6071
1.539 | -83.6042
1.5394 | -83.601
1.5397 | -83.5963
1.54 | -83.5921
1.5404 | -83.5887
1.5407 | -83.5839
1.541 | -83.5782
1.5414 | -83.5748
1.5417 | -83.5722
1.542 | -83.5696
1.5423 | -83.5677
1.5427 | -83.5638
1.543 | -83.562
1.5433 | -83.5607
1.5437 | -83.5585
1.544 | -83.5544
1.5443 | -83.551
1.5446 | -83.5484
1.545 | -83.5451
1.5453 | -83.54
1.5456 | -83.5368
1.546 | -83.5343
1.5463 | -83.5326
1.5466 | -83.5318
1.5469 | -83.5283
1.5473 | -83.5229
1.5476 | -83.5238
1.5479 | -83.5216
1.5483 | -83.5218
1.5486 | -83.52
1.5489 | -83.5194
1.5493 | -83.5214
1.5496 | -83.5204
1.5499 | -83.5182
1.5502 | -83.5176
1.5506 | -83.5192
1.5509 | -83.5176
1.5512 | -83.5147
1.5516 | -83.5119
1.5519 | -83.5095
1.5522 | -83.5095
1.5525 | -83.5064
1.5529 | -83.5044
1.5532 | -83.5021
1.5535 | -83.5028
1.5539 | -83.5048
1.5542 | -83.5058
1.5545 | -83.5021
1.5548 | -83.4977
1.5552 | -83.4984
1.5555 | -83.5
1.5558 | -83.5026
1.5562 | -83.5029
1.5565 | -83.5024
1.5568 | -83.5006
1.5572 | -83.5047
1.5575 | -83.5025
1.5578 | -83.4993
1.5581 | -83.4988
1.5585 | -83.4996
1.5588 | -83.4984
1.5591 | -83.4996
1.5595 | -83.5018
1.5598 | -83.5033
1.5601 | -83.5036
1.5604 | -83.5049
1.5608 | -83.5029
1.5611 | -83.498
1.5614 | -83.4999
1.5618 | -83.501
1.5621 | -83.5012
1.5624 | -83.4978
1.5627 | -83.4955
1.5631 | -83.4971
1.5634 | -83.4973
1.5637 | -83.4992
1.5641 | -83.5003
1.5644 | -83.5
1.5647 | -83.5003
1.5651 | -83.5015
1.5654 | -83.5013
1.5657 | -83.5032
1.566 | -83.5074
1.5664 | -83.5079
1.5667 | -83.5122
1.567 | -83.5142
1.5674 | -83.5163
1.5677 | -83.5177
1.568 | -83.5168
1.5683 | -83.5136
1.5687 | -83.5142
1.569 | -83.5161
1.5693 | -83.5186
1.5697 | -83.5242
1.57 | -83.5305
1.5703 | -83.534
1.5706 | -83.5369
1.571 | -83.5379
1.5713 | -83.5412
1.5716 | -83.542
1.572 | -83.5434
1.5723 | -83.5456
1.5726 | -83.5484
1.573 | -83.5519
1.5733 | -83.5542
1.5736 | -83.5559
1.5739 | -83.5554
1.5743 | -83.5527
1.5746 | -83.5542
1.5749 | -83.5552
1.5753 | -83.5533
1.5756 | -83.5501
1.5759 | -83.5459
1.5762 | -83.5472
1.5766 | -83.5458
1.5769 | -83.5431
1.5772 | -83.5431
1.5776 | -83.5448
1.5779 | -83.5443
1.5782 | -83.5455
1.5785 | -83.5431
1.5789 | -83.5411
1.5792 | -83.5402
1.5795 | -83.5405
1.5799 | -83.5378
1.5802 | -83.5385
1.5805 | -83.5407
1.5808 | -83.5416
1.5812 | -83.5412
1.5815 | -83.5415
1.5818 | -83.5429
1.5822 | -83.5428
1.5825 | -83.5441
1.5828 | -83.544
1.5832 | -83.5437
1.5835 | -83.5417
1.5838 | -83.5427
1.5841 | -83.543
1.5845 | -83.5383
1.5848 | -83.5371
1.5851 | -83.5376
1.5855 | -83.5396
1.5858 | -83.5384
1.5861 | -83.5397
1.5864 | -83.5442
1.5868 | -83.5504
1.5871 | -83.5497
1.5874 | -83.5462
1.5878 | -83.5445
1.5881 | -83.5452
1.5884 | -83.5454
1.5887 | -83.5398
1.5891 | -83.5343
1.5894 | -83.5287
1.5897 | -83.5272
1.5901 | -83.525
1.5904 | -83.5218
1.5907 | -83.5172
1.5911 | -83.5137
1.5914 | -83.513
1.5917 | -83.5099
1.592 | -83.5072
1.5924 | -83.5044
1.5927 | -83.5025
1.593 | -83.4978
1.5934 | -83.4879
1.5937 | -83.481
1.594 | -83.474
1.5943 | -83.4649
1.5947 | -83.4589
1.595 | -83.4563
1.5953 | -83.4494
1.5957 | -83.4445
1.596 | -83.4416
1.5963 | -83.4373
1.5966 | -83.4308
1.597 | -83.4294
1.5973 | -83.4261
1.5976 | -83.4201
1.598 | -83.4122
1.5983 | -83.4049
1.5986 | -83.3981
1.599 | -83.3916
1.5993 | -83.3867
1.5996 | -83.3827
1.5999 | -83.3764
1.6003 | -83.3711
1.6006 | -83.364
1.6009 | -83.356
1.6013 | -83.3469
1.6016 | -83.3397
1.6019 | -83.3324
1.6022 | -83.3251
1.6026 | -83.3192
1.6029 | -83.3114
1.6032 | -83.3041
1.6036 | -83.2978
1.6039 | -83.294
1.6042 | -83.2882
1.6045 | -83.2803
1.6049 | -83.269
1.6052 | -83.2619
1.6055 | -83.2563
1.6059 | -83.2518
1.6062 | -83.2448
1.6065 | -83.2377
1.6069 | -83.2291
1.6072 | -83.2208
1.6075 | -83.2147
1.6078 | -83.2076
1.6082 | -83.2011
1.6085 | -83.1976
1.6088 | -83.1924
1.6092 | -83.1834
1.6095 | -83.1691
1.6098 | -83.1604
1.6101 | -83.1529
1.6105 | -83.1466
1.6108 | -83.1414
1.6111 | -83.1348
1.6115 | -83.1303
1.6118 | -83.1205
1.6121 | -83.1064
1.6124 | -83.0955
1.6128 | -83.0879
1.6131 | -83.081
1.6134 | -83.0738
1.6138 | -83.065
1.6141 | -83.0576
1.6144 | -83.0495
1.6148 | -83.0425
1.6151 | -83.0346
1.6154 | -83.0263
1.6157 | -83.0176
1.6161 | -83.0106
1.6164 | -83.0024
1.6167 | -82.9961
1.6171 | -82.9919
1.6174 | -82.9893
1.6177 | -82.9888
1.618 | -82.9865
1.6184 | -82.9783
1.6187 | -82.97
1.619 | -82.9662
1.6194 | -82.9604
1.6197 | -82.9514
1.62 | -82.9356
1.6203 | -82.9163
1.6207 | -82.9074
1.621 | -82.9021
1.6213 | -82.8981
1.6217 | -82.8917
1.622 | -82.8851
1.6223 | -82.8802
1.6226 | -82.8722
1.623 | -82.8645
1.6233 | -82.8594
1.6236 | -82.8557
1.624 | -82.8482
1.6243 | -82.8399
1.6246 | -82.8347
1.625 | -82.8263
1.6253 | -82.817
1.6256 | -82.8097
1.6259 | -82.8033
1.6263 | -82.7965
1.6266 | -82.7891
1.6269 | -82.7811
1.6273 | -82.7731
1.6276 | -82.7658
1.6279 | -82.7605
1.6282 | -82.7551
1.6286 | -82.7525
1.6289 | -82.7486
1.6292 | -82.7427
1.6296 | -82.7369
1.6299 | -82.7297
1.6302 | -82.7225
1.6305 | -82.7171
1.6309 | -82.7123
1.6312 | -82.7065
1.6315 | -82.7013
1.6319 | -82.6963
1.6322 | -82.6878
1.6325 | -82.679
1.6329 | -82.6712
1.6332 | -82.6666
1.6335 | -82.6602
1.6338 | -82.6528
1.6342 | -82.6486
1.6345 | -82.6465
1.6348 | -82.6428
1.6352 | -82.6384
1.6355 | -82.6361
1.6358 | -82.6337
1.6361 | -82.6309
1.6365 | -82.6268
1.6368 | -82.6231
1.6371 | -82.6185
1.6375 | -82.6146
1.6378 | -82.6093
1.6381 | -82.606
1.6384 | -82.602
1.6388 | -82.5992
1.6391 | -82.596
1.6394 | -82.5922
1.6398 | -82.5886
1.6401 | -82.5849
1.6404 | -82.5807
1.6408 | -82.5779
1.6411 | -82.5755
1.6414 | -82.5733
1.6417 | -82.5737
1.6421 | -82.5757
1.6424 | -82.5748
1.6427 | -82.5729
1.6431 | -82.5712
1.6434 | -82.5711
1.6437 | -82.5719
1.644 | -82.5728
1.6444 | -82.5735
1.6447 | -82.5713
1.645 | -82.5658
1.6454 | -82.5618
1.6457 | -82.5598
1.646 | -82.5579
1.6463 | -82.5554
1.6467 | -82.5547
1.647 | -82.5537
1.6473 | -82.5518
1.6477 | -82.549
1.648 | -82.5437
1.6483 | -82.5388
1.6487 | -82.5361
1.649 | -82.5342
1.6493 | -82.5352
1.6496 | -82.5366
1.65 | -82.5358
1.6503 | -82.5349
1.6506 | -82.5339
1.651 | -82.5307
1.6513 | -82.5264
1.6516 | -82.5247
1.6519 | -82.5242
1.6523 | -82.5243
1.6526 | -82.5237
1.6529 | -82.5219
1.6533 | -82.5185
1.6536 | -82.5163
1.6539 | -82.5141
1.6542 | -82.511
1.6546 | -82.5103
1.6549 | -82.5086
1.6552 | -82.5075
1.6556 | -82.5078
1.6559 | -82.5079
1.6562 | -82.5067
1.6566 | -82.5064
1.6569 | -82.5078
1.6572 | -82.5062
1.6575 | -82.5057
1.6579 | -82.5055
1.6582 | -82.5045
1.6585 | -82.5057
1.6589 | -82.5087
1.6592 | -82.5093
1.6595 | -82.5104
1.6598 | -82.5113
1.6602 | -82.5108
1.6605 | -82.5113
1.6608 | -82.5124
1.6612 | -82.5119
1.6615 | -82.5106
1.6618 | -82.5137
1.6621 | -82.5122
1.6625 | -82.5078
1.6628 | -82.5056
1.6631 | -82.5063
1.6635 | -82.5076
1.6638 | -82.507
1.6641 | -82.5051
1.6644 | -82.5049
1.6648 | -82.5071
1.6651 | -82.5084
1.6654 | -82.509
1.6658 | -82.5084
1.6661 | -82.5095
1.6664 | -82.5098
1.6668 | -82.5117
1.6671 | -82.5146
1.6674 | -82.5155
1.6677 | -82.5158
1.6681 | -82.5152
1.6684 | -82.5162
1.6687 | -82.5157
1.6691 | -82.5151
1.6694 | -82.5147
1.6697 | -82.5155
1.67 | -82.5159
1.6704 | -82.5153
1.6707 | -82.5158
1.671 | -82.5176
1.6714 | -82.5203
1.6717 | -82.5227
1.672 | -82.5234
1.6724 | -82.5245
1.6727 | -82.5258
1.673 | -82.5274
1.6733 | -82.5288
1.6737 | -82.5316
1.674 | -82.5333
1.6743 | -82.5334
1.6747 | -82.5345
1.675 | -82.5375
1.6753 | -82.5421
1.6756 | -82.5446
1.676 | -82.5437
1.6763 | -82.5453
1.6766 | -82.5462
1.677 | -82.5467
1.6773 | -82.5502
1.6776 | -82.5526
1.6779 | -82.5507
1.6783 | -82.548
1.6786 | -82.5463
1.6789 | -82.5456
1.6793 | -82.5482
1.6796 | -82.5482
1.6799 | -82.5485
1.6802 | -82.5491
1.6806 | -82.5526
1.6809 | -82.5539
1.6812 | -82.5562
1.6816 | -82.5611
1.6819 | -82.5671
1.6822 | -82.5688
1.6826 | -82.5697
1.6829 | -82.5705
1.6832 | -82.57
1.6835 | -82.5711
1.6839 | -82.5728
1.6842 | -82.5746
1.6845 | -82.5767
1.6849 | -82.5785
1.6852 | -82.5801
1.6855 | -82.5822
1.6858 | -82.584
1.6862 | -82.5849
1.6865 | -82.5864
1.6868 | -82.5874
1.6872 | -82.59
1.6875 | -82.5959
1.6878 | -82.6019
1.6881 | -82.6053
1.6885 | -82.6072
1.6888 | -82.6097
1.6891 | -82.6163
1.6895 | -82.6207
1.6898 | -82.6224
1.6901 | -82.6265
1.6905 | -82.6268
1.6908 | -82.6283
1.6911 | -82.633
1.6914 | -82.6352
1.6918 | -82.6341
1.6921 | -82.636
1.6924 | -82.6386
1.6928 | -82.6426
1.6931 | -82.6475
1.6934 | -82.6506
1.6937 | -82.6522
1.6941 | -82.6566
1.6944 | -82.6596
1.6947 | -82.6621
1.6951 | -82.6644
1.6954 | -82.6646
1.6957 | -82.6667
1.696 | -82.6668
1.6964 | -82.6653
1.6967 | -82.6636
1.697 | -82.6653
1.6974 | -82.6658
1.6977 | -82.667
1.698 | -82.668
1.6984 | -82.6681
1.6987 | -82.6715
1.699 | -82.6744
1.6993 | -82.6748
1.6997 | -82.6772
1.7 | -82.6793
1.7003 | -82.6806
1.7007 | -82.6807
1.701 | -82.6821
1.7013 | -82.684
1.7016 | -82.6853
1.702 | -82.6869
1.7023 | -82.6907
1.7026 | -82.6927
1.703 | -82.6946
1.7033 | -82.6959
1.7036 | -82.701
1.7039 | -82.7048
1.7043 | -82.7089
1.7046 | -82.7111
1.7049 | -82.713
1.7053 | -82.7149
1.7056 | -82.7161
1.7059 | -82.7195
1.7063 | -82.7216
1.7066 | -82.7236
1.7069 | -82.7258
1.7072 | -82.728
1.7076 | -82.7304
1.7079 | -82.736
1.7082 | -82.7435
1.7086 | -82.7482
1.7089 | -82.7527
1.7092 | -82.7562
1.7095 | -82.7629
1.7099 | -82.7685
1.7102 | -82.7726
1.7105 | -82.7756
1.7109 | -82.7773
1.7112 | -82.7787
1.7115 | -82.7788
1.7118 | -82.782
1.7122 | -82.7863
1.7125 | -82.786
1.7128 | -82.787
1.7132 | -82.7898
1.7135 | -82.7916
1.7138 | -82.7935
1.7142 | -82.7963
1.7145 | -82.7988
1.7148 | -82.8015
1.7151 | -82.8039
1.7155 | -82.8063
1.7158 | -82.8087
1.7161 | -82.8114
1.7165 | -82.8141
1.7168 | -82.8194
1.7171 | -82.8214
1.7174 | -82.823
1.7178 | -82.8261
1.7181 | -82.8288
1.7184 | -82.8317
1.7188 | -82.8379
1.7191 | -82.8413
1.7194 | -82.8452
1.7197 | -82.8488
1.7201 | -82.8519
1.7204 | -82.8548
1.7207 | -82.8596
1.7211 | -82.8624
1.7214 | -82.865
1.7217 | -82.8684
1.7221 | -82.8717
1.7224 | -82.8762
1.7227 | -82.8796
1.723 | -82.8824
1.7234 | -82.887
1.7237 | -82.8906
1.724 | -82.8937
1.7244 | -82.895
1.7247 | -82.8967
1.725 | -82.8996
1.7253 | -82.9029
1.7257 | -82.9094
1.726 | -82.9134
1.7263 | -82.9172
1.7267 | -82.9207
1.727 | -82.9246
1.7273 | -82.9283
1.7276 | -82.9299
1.728 | -82.9349
1.7283 | -82.9409
1.7286 | -82.9512
1.729 | -82.9501
1.7293 | -82.9446
1.7296 | -82.9463
1.7299 | -82.95
1.7303 | -82.9541
1.7306 | -82.9569
1.7309 | -82.9612
1.7313 | -82.9664
1.7316 | -82.969
1.7319 | -82.9725
1.7323 | -82.978
1.7326 | -82.9832
1.7329 | -82.9861
1.7332 | -82.9913
1.7336 | -82.9951
1.7339 | -82.9963
1.7342 | -82.9966
1.7346 | -82.9994
1.7349 | -83.0031
1.7352 | -83.0093
1.7355 | -83.017
1.7359 | -83.0169
1.7362 | -83.0156
1.7365 | -83.0209
1.7369 | -83.0333
1.7372 | -83.0492
1.7375 | -83.0511
1.7378 | -83.0575
1.7382 | -83.0614
1.7385 | -83.0653
1.7388 | -83.0695
1.7392 | -83.0725
1.7395 | -83.0727
1.7398 | -83.0737
1.7402 | -83.0779
1.7405 | -83.0781
1.7408 | -83.0765
1.7411 | -83.0763
1.7415 | -83.0794
1.7418 | -83.0832
1.7421 | -83.0862
1.7425 | -83.0933
1.7428 | -83.0999
1.7431 | -83.1028
1.7434 | -83.1056
1.7438 | -83.1085
1.7441 | -83.1096
1.7444 | -83.1108
1.7448 | -83.1152
1.7451 | -83.1187
1.7454 | -83.1194
1.7457 | -83.1213
1.7461 | -83.1215
1.7464 | -83.1176
1.7467 | -83.1193
1.7471 | -83.1172
1.7474 | -83.1192
1.7477 | -83.1238
1.7481 | -83.1274
1.7484 | -83.1292
1.7487 | -83.1308
1.749 | -83.1252
1.7494 | -83.1234
1.7497 | -83.1251
1.75 | -83.1287
1.7504 | -83.1339
1.7507 | -83.137
1.751 | -83.1372
1.7513 | -83.1324
1.7517 | -83.131
1.752 | -83.1326
1.7523 | -83.1338
1.7527 | -83.1337
1.753 | -83.1335
1.7533 | -83.1332
1.7536 | -83.132
1.754 | -83.1312
1.7543 | -83.1298
1.7546 | -83.128
1.755 | -83.1279
1.7553 | -83.1278
1.7556 | -83.1272
1.756 | -83.1279
1.7563 | -83.124
1.7566 | -83.119
1.7569 | -83.1194
1.7573 | -83.1199
1.7576 | -83.1181
1.7579 | -83.1172
1.7583 | -83.1162
1.7586 | -83.1147
1.7589 | -83.1147
1.7592 | -83.1119
1.7596 | -83.1036
1.7599 | -83.0965
1.7602 | -83.0941
1.7606 | -83.0912
1.7609 | -83.0875
1.7612 | -83.0868
1.7615 | -83.0862
1.7619 | -83.0835
1.7622 | -83.0807
1.7625 | -83.0767
1.7629 | -83.0728
1.7632 | -83.0709
1.7635 | -83.0694
1.7638 | -83.0668
1.7642 | -83.0642
1.7645 | -83.061
1.7648 | -83.0574
1.7652 | -83.0525
1.7655 | -83.0459
1.7658 | -83.0414
1.7662 | -83.0301
1.7665 | -83.0213
1.7668 | -83.0159
1.7671 | -83.0123
1.7675 | -83.0109
1.7678 | -83.007
1.7681 | -83.0025
1.7685 | -82.9971
1.7688 | -82.9899
1.7691 | -82.984
1.7694 | -82.978
1.7698 | -82.9719
1.7701 | -82.9701
1.7704 | -82.9687
1.7708 | -82.9636
1.7711 | -82.9577
1.7714 | -82.9514
1.7717 | -82.946
1.7721 | -82.9397
1.7724 | -82.9352
1.7727 | -82.9301
1.7731 | -82.9246
1.7734 | -82.9206
1.7737 | -82.9152
1.7741 | -82.9086
1.7744 | -82.9034
1.7747 | -82.8945
1.775 | -82.8867
1.7754 | -82.8803
1.7757 | -82.873
1.776 | -82.8676
1.7764 | -82.8627
1.7767 | -82.8557
1.777 | -82.8487
1.7773 | -82.8423
1.7777 | -82.8361
1.778 | -82.8303
1.7783 | -82.8221
1.7787 | -82.815
1.779 | -82.8082
1.7793 | -82.8002
1.7796 | -82.794
1.78 | -82.7875
1.7803 | -82.7799
1.7806 | -82.7715
1.781 | -82.7658
1.7813 | -82.7627
1.7816 | -82.7569
1.782 | -82.7511
1.7823 | -82.7475
1.7826 | -82.7454
1.7829 | -82.7424
1.7833 | -82.7371
1.7836 | -82.7326
1.7839 | -82.7297
1.7843 | -82.727
1.7846 | -82.7211
1.7849 | -82.7144
1.7852 | -82.7124
1.7856 | -82.7068
1.7859 | -82.7033
1.7862 | -82.7001
1.7866 | -82.6948
1.7869 | -82.6857
1.7872 | -82.6792
1.7875 | -82.6736
1.7879 | -82.6674
1.7882 | -82.6633
1.7885 | -82.6587
1.7889 | -82.6533
1.7892 | -82.6502
1.7895 | -82.6477
1.7899 | -82.6435
1.7902 | -82.6389
1.7905 | -82.6367
1.7908 | -82.6351
1.7912 | -82.6306
1.7915 | -82.6256
1.7918 | -82.6218
1.7922 | -82.6145
1.7925 | -82.6076
1.7928 | -82.6026
1.7931 | -82.5982
1.7935 | -82.5945
1.7938 | -82.5906
1.7941 | -82.5863
1.7945 | -82.5836
1.7948 | -82.5857
1.7951 | -82.5864
1.7954 | -82.5852
1.7958 | -82.5823
1.7961 | -82.5784
1.7964 | -82.577
1.7968 | -82.5759
1.7971 | -82.5756
1.7974 | -82.5738
1.7978 | -82.5715
1.7981 | -82.5701
1.7984 | -82.5692
1.7987 | -82.5689
1.7991 | -82.5706
1.7994 | -82.5746
1.7997 | -82.5769
1.8001 | -82.5783
1.8004 | -82.5773
1.8007 | -82.5783
1.801 | -82.5789
1.8014 | -82.5769
1.8017 | -82.5719
1.802 | -82.5717
1.8024 | -82.5706
1.8027 | -82.5701
1.803 | -82.5704
1.8033 | -82.57
1.8037 | -82.5698
1.804 | -82.5755
1.8043 | -82.5768
1.8047 | -82.577
1.805 | -82.5795
1.8053 | -82.5826
1.8057 | -82.5844
1.806 | -82.5857
1.8063 | -82.5869
1.8066 | -82.588
1.807 | -82.5888
1.8073 | -82.5858
1.8076 | -82.5857
1.808 | -82.5891
1.8083 | -82.5924
1.8086 | -82.596
1.8089 | -82.6002
1.8093 | -82.6039
1.8096 | -82.606
1.8099 | -82.6053
1.8103 | -82.6042
1.8106 | -82.6084
1.8109 | -82.6121
1.8112 | -82.6172
1.8116 | -82.6212
1.8119 | -82.6247
1.8122 | -82.6255
1.8126 | -82.6266
1.8129 | -82.6325
1.8132 | -82.6362
1.8135 | -82.6441
1.8139 | -82.648
1.8142 | -82.6535
1.8145 | -82.6592
1.8149 | -82.6626
1.8152 | -82.6671
1.8155 | -82.6718
1.8159 | -82.6769
1.8162 | -82.6819
1.8165 | -82.6872
1.8168 | -82.6929
1.8172 | -82.6985
1.8175 | -82.7043
1.8178 | -82.7111
1.8182 | -82.7153
1.8185 | -82.7206
1.8188 | -82.7289
1.8191 | -82.7383
1.8195 | -82.7435
1.8198 | -82.7502
1.8201 | -82.7581
1.8205 | -82.7659
1.8208 | -82.7737
1.8211 | -82.7803
1.8214 | -82.7866
1.8218 | -82.7936
1.8221 | -82.8013
1.8224 | -82.8069
1.8228 | -82.8099
1.8231 | -82.8184
1.8234 | -82.8253
1.8238 | -82.834
1.8241 | -82.8408
1.8244 | -82.8422
1.8247 | -82.8478
1.8251 | -82.8544
1.8254 | -82.8609
1.8257 | -82.8683
1.8261 | -82.8755
1.8264 | -82.8829
1.8267 | -82.8895
1.827 | -82.8961
1.8274 | -82.9017
1.8277 | -82.9094
1.828 | -82.9169
1.8284 | -82.9234
1.8287 | -82.9289
1.829 | -82.933
1.8293 | -82.9403
1.8297 | -82.9478
1.83 | -82.955
1.8303 | -82.9622
1.8307 | -82.9689
1.831 | -82.9735
1.8313 | -82.9787
1.8317 | -82.9862
1.832 | -82.9938
1.8323 | -83.0014
1.8326 | -83.0072
1.833 | -83.0124
1.8333 | -83.017
1.8336 | -83.0232
1.834 | -83.0298
1.8343 | -83.036
1.8346 | -83.0417
1.8349 | -83.0483
1.8353 | -83.0562
1.8356 | -83.0627
1.8359 | -83.0692
1.8363 | -83.0787
1.8366 | -83.0842
1.8369 | -83.0845
1.8372 | -83.0865
1.8376 | -83.0918
1.8379 | -83.0975
1.8382 | -83.1038
1.8386 | -83.1101
1.8389 | -83.1147
1.8392 | -83.1202
1.8396 | -83.1254
1.8399 | -83.1314
1.8402 | -83.138
1.8405 | -83.1419
1.8409 | -83.1484
1.8412 | -83.1553
1.8415 | -83.1605
1.8419 | -83.1652
1.8422 | -83.1649
1.8425 | -83.1703
1.8428 | -83.1749
1.8432 | -83.1792
1.8435 | -83.1839
1.8438 | -83.1877
1.8442 | -83.1917
1.8445 | -83.1953
1.8448 | -83.2009
1.8451 | -83.2049
1.8455 | -83.2093
1.8458 | -83.2133
1.8461 | -83.216
1.8465 | -83.2196
1.8468 | -83.2234
1.8471 | -83.2276
1.8475 | -83.2329
1.8478 | -83.2364
1.8481 | -83.2398
1.8484 | -83.243
1.8488 | -83.2487
1.8491 | -83.2536
1.8494 | -83.2544
1.8498 | -83.2556
1.8501 | -83.2609
1.8504 | -83.2643
1.8507 | -83.2665
1.8511 | -83.269
1.8514 | -83.2713
1.8517 | -83.2702
1.8521 | -83.2714
1.8524 | -83.2751
1.8527 | -83.2799
1.853 | -83.2819
1.8534 | -83.2846
1.8537 | -83.2861
1.854 | -83.2934
1.8544 | -83.3022
1.8547 | -83.3044
1.855 | -83.302
1.8554 | -83.3023
1.8557 | -83.303
1.856 | -83.3023
1.8563 | -83.3022
1.8567 | -83.3034
1.857 | -83.3029
1.8573 | -83.3026
1.8577 | -83.2999
1.858 | -83.2995
1.8583 | -83.2993
1.8586 | -83.2993
1.859 | -83.3014
1.8593 | -83.3034
1.8596 | -83.3029
1.86 | -83.2987
1.8603 | -83.2949
1.8606 | -83.2895
1.8609 | -83.289
1.8613 | -83.2855
1.8616 | -83.2886
1.8619 | -83.2856
1.8623 | -83.2815
1.8626 | -83.2783
1.8629 | -83.2764
1.8633 | -83.2736
1.8636 | -83.2695
1.8639 | -83.2635
1.8642 | -83.2586
1.8646 | -83.2553
1.8649 | -83.2537
1.8652 | -83.2539
1.8656 | -83.2517
1.8659 | -83.2445
1.8662 | -83.2429
1.8665 | -83.2441
1.8669 | -83.2447
1.8672 | -83.2413
1.8675 | -83.239
1.8679 | -83.2352
1.8682 | -83.2318
1.8685 | -83.2284
1.8688 | -83.2252
1.8692 | -83.2222
1.8695 | -83.2184
1.8698 | -83.2164
1.8702 | -83.2142
1.8705 | -83.2066
1.8708 | -83.2035
1.8711 | -83.2002
1.8715 | -83.1987
1.8718 | -83.1954
1.8721 | -83.1922
1.8725 | -83.1938
1.8728 | -83.1903
1.8731 | -83.1864
1.8735 | -83.1837
1.8738 | -83.1805
1.8741 | -83.1762
1.8744 | -83.1724
1.8748 | -83.1684
1.8751 | -83.1627
1.8754 | -83.1636
1.8758 | -83.1621
1.8761 | -83.1583
1.8764 | -83.1532
1.8767 | -83.1489
1.8771 | -83.1449
1.8774 | -83.1391
1.8777 | -83.1401
1.8781 | -83.1384
1.8784 | -83.1329
1.8787 | -83.1263
1.879 | -83.1221
1.8794 | -83.1173
1.8797 | -83.1118
1.88 | -83.1065
1.8804 | -83.103
1.8807 | -83.0995
1.881 | -83.0966
1.8814 | -83.094
1.8817 | -83.0902
1.882 | -83.0852
1.8823 | -83.0805
1.8827 | -83.0759
1.883 | -83.0724
1.8833 | -83.0698
1.8837 | -83.0664
1.884 | -83.0638
1.8843 | -83.0602
1.8846 | -83.0572
1.885 | -83.054
1.8853 | -83.0506
1.8856 | -83.0469
1.886 | -83.0419
1.8863 | -83.0394
1.8866 | -83.0415
1.8869 | -83.0426
1.8873 | -83.0396
1.8876 | -83.036
1.8879 | -83.0342
1.8883 | -83.0307
1.8886 | -83.0273
1.8889 | -83.0234
1.8893 | -83.0178
1.8896 | -83.0128
1.8899 | -83.0096
1.8902 | -83.0036
1.8906 | -82.9963
1.8909 | -82.9933
1.8912 | -82.9922
1.8916 | -82.988
1.8919 | -82.9832
1.8922 | -82.9768
1.8925 | -82.9687
1.8929 | -82.9647
1.8932 | -82.9619
1.8935 | -82.9588
1.8939 | -82.9605
1.8942 | -82.962
1.8945 | -82.9614
1.8948 | -82.9608
1.8952 | -82.9568
1.8955 | -82.9532
1.8958 | -82.95
1.8962 | -82.9475
1.8965 | -82.9466
1.8968 | -82.9429
1.8972 | -82.9403
1.8975 | -82.9393
1.8978 | -82.9355
1.8981 | -82.933
1.8985 | -82.9291
1.8988 | -82.9255
1.8991 | -82.9258
1.8995 | -82.9238
1.8998 | -82.9225
1.9001 | -82.9208
1.9004 | -82.917
1.9008 | -82.9156
1.9011 | -82.9139
1.9014 | -82.9108
1.9018 | -82.9085
1.9021 | -82.9065
1.9024 | -82.9042
1.9027 | -82.9048
1.9031 | -82.9041
1.9034 | -82.8977
1.9037 | -82.8917
1.9041 | -82.8857
1.9044 | -82.8789
1.9047 | -82.8757
1.905 | -82.874
1.9054 | -82.8728
1.9057 | -82.8711
1.906 | -82.8686
1.9064 | -82.8671
1.9067 | -82.8663
1.907 | -82.8639
1.9074 | -82.8595
1.9077 | -82.8567
1.908 | -82.8541
1.9083 | -82.8532
1.9087 | -82.8504
1.909 | -82.8479
1.9093 | -82.8467
1.9097 | -82.8435
1.91 | -82.8395
1.9103 | -82.8372
1.9106 | -82.8354
1.911 | -82.833
1.9113 | -82.8296
1.9116 | -82.8245
1.912 | -82.8221
1.9123 | -82.8203
1.9126 | -82.8184
1.9129 | -82.8157
1.9133 | -82.8132
1.9136 | -82.8124
1.9139 | -82.8102
1.9143 | -82.8078
1.9146 | -82.8048
1.9149 | -82.8001
1.9153 | -82.7977
1.9156 | -82.796
1.9159 | -82.797
1.9162 | -82.7974
1.9166 | -82.7951
1.9169 | -82.7895
1.9172 | -82.7877
1.9176 | -82.7867
1.9179 | -82.7848
1.9182 | -82.7839
1.9185 | -82.7821
1.9189 | -82.7797
1.9192 | -82.7783
1.9195 | -82.7771
1.9199 | -82.7749
1.9202 | -82.7727
1.9205 | -82.772
1.9208 | -82.7706
1.9212 | -82.77
1.9215 | -82.7714
1.9218 | -82.7719
1.9222 | -82.7699
1.9225 | -82.7673
1.9228 | -82.761
1.9232 | -82.7572
1.9235 | -82.7561
1.9238 | -82.7547
1.9241 | -82.7532
1.9245 | -82.7484
1.9248 | -82.7445
1.9251 | -82.7417
1.9255 | -82.7402
1.9258 | -82.7396
1.9261 | -82.7348
1.9264 | -82.7316
1.9268 | -82.7297
1.9271 | -82.7283
1.9274 | -82.7272
1.9278 | -82.7225
1.9281 | -82.7211
1.9284 | -82.7167
1.9287 | -82.7095
1.9291 | -82.7037
1.9294 | -82.7002
1.9297 | -82.6983
1.9301 | -82.6957
1.9304 | -82.6922
1.9307 | -82.6905
1.9311 | -82.6881
1.9314 | -82.6875
1.9317 | -82.6886
1.932 | -82.6833
1.9324 | -82.6801
1.9327 | -82.6773
1.933 | -82.6745
1.9334 | -82.6734
1.9337 | -82.6733
1.934 | -82.6709
1.9343 | -82.666
1.9347 | -82.6616
1.935 | -82.6587
1.9353 | -82.6565
1.9357 | -82.6539
1.936 | -82.6519
1.9363 | -82.6502
1.9366 | -82.6455
1.937 | -82.6382
1.9373 | -82.6354
1.9376 | -82.6336
1.938 | -82.6361
1.9383 | -82.6361
1.9386 | -82.6332
1.939 | -82.6296
1.9393 | -82.6265
1.9396 | -82.6244
1.9399 | -82.6218
1.9403 | -82.6208
1.9406 | -82.6221
1.9409 | -82.6239
1.9413 | -82.6204
1.9416 | -82.6157
1.9419 | -82.6122
1.9422 | -82.6088
1.9426 | -82.6049
1.9429 | -82.6
1.9432 | -82.5965
1.9436 | -82.5926
1.9439 | -82.5888
1.9442 | -82.5864
1.9445 | -82.5838
1.9449 | -82.5813
1.9452 | -82.5792
1.9455 | -82.5723
1.9459 | -82.5682
1.9462 | -82.5652
1.9465 | -82.5576
1.9469 | -82.5542
1.9472 | -82.5499
1.9475 | -82.5441
1.9478 | -82.5409
1.9482 | -82.5368
1.9485 | -82.5333
1.9488 | -82.5294
1.9492 | -82.5255
1.9495 | -82.5215
1.9498 | -82.5156
1.9501 | -82.5105
1.9505 | -82.5074
1.9508 | -82.5055
1.9511 | -82.5033
1.9515 | -82.4987
1.9518 | -82.4924
1.9521 | -82.4886
1.9524 | -82.4857
1.9528 | -82.4797
1.9531 | -82.4726
1.9534 | -82.4667
1.9538 | -82.4635
1.9541 | -82.4621
1.9544 | -82.459
1.9547 | -82.4555
1.9551 | -82.4518
1.9554 | -82.4479
1.9557 | -82.4452
1.9561 | -82.4404
1.9564 | -82.44
1.9567 | -82.4364
1.9571 | -82.4315
1.9574 | -82.4271
1.9577 | -82.4223
1.958 | -82.4173
1.9584 | -82.4139
1.9587 | -82.4082
1.959 | -82.4022
1.9594 | -82.3985
1.9597 | -82.3939
1.96 | -82.3878
1.9603 | -82.3804
1.9607 | -82.3785
1.961 | -82.3774
1.9613 | -82.3708
1.9617 | -82.3646
1.962 | -82.3581
1.9623 | -82.3484
1.9626 | -82.3403
1.963 | -82.333
1.9633 | -82.3241
1.9636 | -82.318
1.964 | -82.3131
1.9643 | -82.3074
1.9646 | -82.3036
1.965 | -82.2985
1.9653 | -82.2946
1.9656 | -82.2853
1.9659 | -82.2819
1.9663 | -82.2757
1.9666 | -82.2701
1.9669 | -82.2655
1.9673 | -82.262
1.9676 | -82.2597
1.9679 | -82.2534
1.9682 | -82.2475
1.9686 | -82.2421
1.9689 | -82.2362
1.9692 | -82.2305
1.9696 | -82.228
1.9699 | -82.2277
1.9702 | -82.2243
1.9705 | -82.2193
1.9709 | -82.2137
1.9712 | -82.2077
1.9715 | -82.2025
1.9719 | -82.1967
1.9722 | -82.1916
1.9725 | -82.1862
1.9729 | -82.1798
1.9732 | -82.1735
1.9735 | -82.1667
1.9738 | -82.161
1.9742 | -82.1527
1.9745 | -82.1445
1.9748 | -82.1399
1.9752 | -82.1353
1.9755 | -82.1304
1.9758 | -82.1213
1.9761 | -82.112
1.9765 | -82.1055
1.9768 | -82.0991
1.9771 | -82.0936
1.9775 | -82.0883
1.9778 | -82.0815
1.9781 | -82.0756
1.9784 | -82.0696
1.9788 | -82.065
1.9791 | -82.0564
1.9794 | -82.0462
1.9798 | -82.0394
1.9801 | -82.0328
1.9804 | -82.0275
1.9808 | -82.0211
1.9811 | -82.0138
1.9814 | -82.0082
1.9817 | -82.0031
1.9821 | -81.996
1.9824 | -81.99
1.9827 | -81.9864
1.9831 | -81.9825
1.9834 | -81.977
1.9837 | -81.971
1.984 | -81.9648
1.9844 | -81.9585
1.9847 | -81.9494
1.985 | -81.9435
1.9854 | -81.9375
1.9857 | -81.9323
1.986 | -81.9264
1.9863 | -81.9212
1.9867 | -81.9159
1.987 | -81.9095
1.9873 | -81.9028
1.9877 | -81.897
1.988 | -81.8932
1.9883 | -81.8902
1.9887 | -81.885
1.989 | -81.8816
1.9893 | -81.8767
1.9896 | -81.8728
1.99 | -81.8682
1.9903 | -81.8622
1.9906 | -81.8561
1.991 | -81.8516
1.9913 | -81.846
1.9916 | -81.8403
1.9919 | -81.8331
1.9923 | -81.8263
1.9926 | -81.8201
1.9929 | -81.8143
1.9933 | -81.8083
1.9936 | -81.802
1.9939 | -81.7955
1.9942 | -81.7887
1.9946 | -81.7859
1.9949 | -81.7827
1.9952 | -81.7791
1.9956 | -81.7744
1.9959 | -81.7668
1.9962 | -81.7608
1.9966 | -81.7567
1.9969 | -81.7532
1.9972 | -81.7476
1.9975 | -81.7416
1.9979 | -81.7385
1.9982 | -81.7346
1.9985 | -81.7306
1.9989 | -81.7267
1.9992 | -81.7215
1.9995 | -81.7167
1.9998 | -81.7125
2.0002 | -81.7093
2.0005 | -81.7064
2.0008 | -81.7008
2.0012 | -81.6955
2.0015 | -81.6899
2.0018 | -81.6837
2.0021 | -81.6771
2.0025 | -81.6728
2.0028 | -81.6678
2.0031 | -81.6633
2.0035 | -81.6602
2.0038 | -81.6543
2.0041 | -81.6499
2.0044 | -81.6464
2.0048 | -81.6423
2.0051 | -81.6372
2.0054 | -81.6331
2.0058 | -81.6297
2.0061 | -81.6258
2.0064 | -81.6219
2.0068 | -81.6182
2.0071 | -81.6148
2.0074 | -81.6117
2.0077 | -81.6066
2.0081 | -81.6012
2.0084 | -81.5975
2.0087 | -81.5953
2.0091 | -81.5907
2.0094 | -81.5855
2.0097 | -81.5829
2.01 | -81.5806
2.0104 | -81.5792
2.0107 | -81.5779
2.011 | -81.5753
2.0114 | -81.5697
2.0117 | -81.5673
2.012 | -81.5643
2.0123 | -81.5608
2.0127 | -81.5595
2.013 | -81.5595
2.0133 | -81.5574
2.0137 | -81.5537
2.014 | -81.5519
2.0143 | -81.5502
2.0147 | -81.5476
2.015 | -81.5453
2.0153 | -81.5433
2.0156 | -81.5424
2.016 | -81.5408
2.0163 | -81.5385
2.0166 | -81.5355
2.017 | -81.5335
2.0173 | -81.5304
2.0176 | -81.5288
2.0179 | -81.5258
2.0183 | -81.5245
2.0186 | -81.5237
2.0189 | -81.523
2.0193 | -81.5214
2.0196 | -81.5201
2.0199 | -81.5189
2.0202 | -81.5169
2.0206 | -81.5128
2.0209 | -81.5081
2.0212 | -81.5055
2.0216 | -81.5048
2.0219 | -81.5046
2.0222 | -81.5027
2.0226 | -81.4996
2.0229 | -81.4965
2.0232 | -81.4953
2.0235 | -81.4944
2.0239 | -81.4941
2.0242 | -81.4932
2.0245 | -81.4931
2.0249 | -81.4926
2.0252 | -81.4914
2.0255 | -81.492
2.0258 | -81.4938
2.0262 | -81.4949
2.0265 | -81.4947
2.0268 | -81.4954
2.0272 | -81.498
2.0275 | -81.496
2.0278 | -81.4925
2.0281 | -81.4943
2.0285 | -81.4954
2.0288 | -81.4966
2.0291 | -81.5004
2.0295 | -81.5023
2.0298 | -81.5016
2.0301 | -81.5008
2.0305 | -81.5003
2.0308 | -81.5004
2.0311 | -81.4989
2.0314 | -81.4986
2.0318 | -81.4993
2.0321 | -81.5002
2.0324 | -81.5011
2.0328 | -81.5004
2.0331 | -81.4995
2.0334 | -81.4997
2.0337 | -81.5015
2.0341 | -81.5023
2.0344 | -81.5026
2.0347 | -81.5029
2.0351 | -81.5027
2.0354 | -81.5026
2.0357 | -81.5031
2.036 | -81.5033
2.0364 | -81.5043
2.0367 | -81.5058
2.037 | -81.5066
2.0374 | -81.5045
2.0377 | -81.5058
2.038 | -81.5072
2.0384 | -81.5087
2.0387 | -81.5108
2.039 | -81.5138
2.0393 | -81.5149
2.0397 | -81.5155
2.04 | -81.5159
2.0403 | -81.5169
2.0407 | -81.5214
2.041 | -81.5257
2.0413 | -81.5295
2.0416 | -81.5305
2.042 | -81.5349
2.0423 | -81.5351
2.0426 | -81.5363
2.043 | -81.5389
2.0433 | -81.5386
2.0436 | -81.5388
2.0439 | -81.5388
2.0443 | -81.5395
2.0446 | -81.5416
2.0449 | -81.5434
2.0453 | -81.5432
2.0456 | -81.5451
2.0459 | -81.546
2.0463 | -81.5469
2.0466 | -81.5508
2.0469 | -81.5539
2.0472 | -81.5554
2.0476 | -81.5577
2.0479 | -81.5608
2.0482 | -81.5608
2.0486 | -81.5602
2.0489 | -81.5613
2.0492 | -81.5645
2.0495 | -81.5685
2.0499 | -81.5711
2.0502 | -81.5721
2.0505 | -81.5696
2.0509 | -81.5684
2.0512 | -81.562
2.0515 | -81.5614
2.0518 | -81.5614
2.0522 | -81.5621
2.0525 | -81.5635
2.0528 | -81.5658
2.0532 | -81.5692
2.0535 | -81.5703
2.0538 | -81.5719
2.0541 | -81.574
2.0545 | -81.5748
2.0548 | -81.5756
2.0551 | -81.5778
2.0555 | -81.578
2.0558 | -81.5774
2.0561 | -81.5775
2.0565 | -81.5793
2.0568 | -81.58
2.0571 | -81.5814
2.0574 | -81.5837
2.0578 | -81.5853
2.0581 | -81.5853
2.0584 | -81.5867
2.0588 | -81.5894
2.0591 | -81.5903
2.0594 | -81.5911
2.0597 | -81.5915
2.0601 | -81.5921
2.0604 | -81.5944
2.0607 | -81.596
2.0611 | -81.5985
2.0614 | -81.6003
2.0617 | -81.6036
2.062 | -81.6053
2.0624 | -81.6059
2.0627 | -81.6055
2.063 | -81.6046
2.0634 | -81.608
2.0637 | -81.611
2.064 | -81.6127
2.0644 | -81.6133
2.0647 | -81.614
2.065 | -81.6175
2.0653 | -81.6197
2.0657 | -81.6211
2.066 | -81.62
2.0663 | -81.6201
2.0667 | -81.6234
2.067 | -81.6248
2.0673 | -81.6261
2.0676 | -81.6292
2.068 | -81.6342
2.0683 | -81.6384
2.0686 | -81.6399
2.069 | -81.6385
2.0693 | -81.638
2.0696 | -81.6374
2.0699 | -81.6366
2.0703 | -81.6345
2.0706 | -81.6343
2.0709 | -81.6339
2.0713 | -81.635
2.0716 | -81.6382
2.0719 | -81.6397
2.0723 | -81.6386
2.0726 | -81.6367
2.0729 | -81.6336
2.0732 | -81.6328
2.0736 | -81.6345
2.0739 | -81.6372
2.0742 | -81.6381
2.0746 | -81.6393
2.0749 | -81.64
2.0752 | -81.6402
2.0755 | -81.6399
2.0759 | -81.6397
2.0762 | -81.6427
2.0765 | -81.6473
2.0769 | -81.6462
2.0772 | -81.646
2.0775 | -81.6479
2.0778 | -81.6487
2.0782 | -81.648
2.0785 | -81.6483
2.0788 | -81.6491
2.0792 | -81.6489
2.0795 | -81.6491
2.0798 | -81.6489
2.0802 | -81.6478
2.0805 | -81.6451
2.0808 | -81.644
2.0811 | -81.6445
2.0815 | -81.6462
2.0818 | -81.6479
2.0821 | -81.6493
2.0825 | -81.6501
2.0828 | -81.65
2.0831 | -81.6529
2.0834 | -81.6543
2.0838 | -81.6537
2.0841 | -81.6543
2.0844 | -81.6551
2.0848 | -81.6553
2.0851 | -81.6541
2.0854 | -81.6538
2.0857 | -81.657
2.0861 | -81.658
2.0864 | -81.658
2.0867 | -81.6598
2.0871 | -81.6595
2.0874 | -81.6599
2.0877 | -81.6595
2.0881 | -81.6607
2.0884 | -81.6606
2.0887 | -81.6615
2.089 | -81.6628
2.0894 | -81.6622
2.0897 | -81.662
2.09 | -81.6625
2.0904 | -81.6646
2.0907 | -81.6657
2.091 | -81.664
2.0913 | -81.6616
2.0917 | -81.6608
2.092 | -81.6598
2.0923 | -81.6595
2.0927 | -81.6615
2.093 | -81.6625
2.0933 | -81.6633
2.0936 | -81.6657
2.094 | -81.6691
2.0943 | -81.6709
2.0946 | -81.671
2.095 | -81.6712
2.0953 | -81.671
2.0956 | -81.6718
2.096 | -81.6715
2.0963 | -81.6721
2.0966 | -81.6724
2.0969 | -81.6728
2.0973 | -81.6751
2.0976 | -81.6749
2.0979 | -81.6757
2.0983 | -81.6772
2.0986 | -81.6774
2.0989 | -81.6788
2.0992 | -81.6775
2.0996 | -81.6759
2.0999 | -81.676
2.1002 | -81.6773
2.1006 | -81.6791
2.1009 | -81.68
2.1012 | -81.6806
2.1015 | -81.6824
2.1019 | -81.685
2.1022 | -81.6862
2.1025 | -81.6863
2.1029 | -81.6871
2.1032 | -81.6878
2.1035 | -81.6902
2.1038 | -81.6949
2.1042 | -81.6984
2.1045 | -81.7024
2.1048 | -81.7073
2.1052 | -81.7082
2.1055 | -81.7051
2.1058 | -81.7051
2.1062 | -81.7057
2.1065 | -81.7066
2.1068 | -81.7077
2.1071 | -81.7105
2.1075 | -81.7121
2.1078 | -81.7129
2.1081 | -81.7164
2.1085 | -81.7175
2.1088 | -81.7184
2.1091 | -81.7194
2.1094 | -81.7193
2.1098 | -81.7202
2.1101 | -81.722
2.1104 | -81.7215
2.1108 | -81.7222
2.1111 | -81.7233
2.1114 | -81.7226
2.1117 | -81.7214
2.1121 | -81.7206
2.1124 | -81.7215
2.1127 | -81.7226
2.1131 | -81.7239
2.1134 | -81.7254
2.1137 | -81.7261
2.1141 | -81.7265
2.1144 | -81.7266
2.1147 | -81.7272
2.115 | -81.7291
2.1154 | -81.7309
2.1157 | -81.7336
2.116 | -81.7338
2.1164 | -81.7366
2.1167 | -81.7395
2.117 | -81.7408
2.1173 | -81.7396
2.1177 | -81.7408
2.118 | -81.7429
2.1183 | -81.745
2.1187 | -81.7483
2.119 | -81.7509
2.1193 | -81.7527
2.1196 | -81.7546
2.12 | -81.7551
2.1203 | -81.7568
2.1206 | -81.7598
2.121 | -81.7617
2.1213 | -81.7647
2.1216 | -81.7648
2.122 | -81.7616
2.1223 | -81.7646
2.1226 | -81.7653
2.1229 | -81.7649
2.1233 | -81.7655
2.1236 | -81.7676
2.1239 | -81.7706
2.1243 | -81.7722
2.1246 | -81.7729
2.1249 | -81.7749
2.1252 | -81.7765
2.1256 | -81.7765
2.1259 | -81.7756
2.1262 | -81.7771
2.1266 | -81.7784
2.1269 | -81.776
2.1272 | -81.7682
2.1275 | -81.7675
2.1279 | -81.7719
2.1282 | -81.7741
2.1285 | -81.7757
2.1289 | -81.7768
2.1292 | -81.7781
2.1295 | -81.78
2.1299 | -81.7802
2.1302 | -81.78
2.1305 | -81.7813
2.1308 | -81.7832
2.1312 | -81.785
2.1315 | -81.7864
2.1318 | -81.7868
2.1322 | -81.7885
2.1325 | -81.7894
2.1328 | -81.791
2.1331 | -81.7923
2.1335 | -81.7938
2.1338 | -81.7938
2.1341 | -81.7933
2.1345 | -81.7951
2.1348 | -81.7972
2.1351 | -81.8009
2.1354 | -81.8055
2.1358 | -81.8077
2.1361 | -81.8093
2.1364 | -81.8083
2.1368 | -81.8093
2.1371 | -81.8114
2.1374 | -81.8137
2.1378 | -81.8182
2.1381 | -81.8226
2.1384 | -81.8231
2.1387 | -81.8213
2.1391 | -81.8224
2.1394 | -81.8247
2.1397 | -81.826
2.1401 | -81.8292
2.1404 | -81.8307
2.1407 | -81.8312
2.141 | -81.8298
2.1414 | -81.8317
2.1417 | -81.8344
2.142 | -81.8351
2.1424 | -81.8359
2.1427 | -81.8367
2.143 | -81.8371
2.1433 | -81.8365
2.1437 | -81.8364
2.144 | -81.836
2.1443 | -81.8352
2.1447 | -81.8328
2.145 | -81.8305
2.1453 | -81.8307
2.1456 | -81.8315
2.146 | -81.8233
2.1463 | -81.8239
2.1466 | -81.8264
2.147 | -81.8255
2.1473 | -81.8237
2.1476 | -81.8219
2.148 | -81.8207
2.1483 | -81.8207
2.1486 | -81.818
2.1489 | -81.8125
2.1493 | -81.8064
2.1496 | -81.8056
2.1499 | -81.8074
2.1503 | -81.809
2.1506 | -81.8078
2.1509 | -81.8066
2.1512 | -81.8049
2.1516 | -81.8026
2.1519 | -81.8018
2.1522 | -81.8008
2.1526 | -81.8007
2.1529 | -81.8012
2.1532 | -81.8009
2.1535 | -81.7982
2.1539 | -81.7972
2.1542 | -81.7985
2.1545 | -81.7967
2.1549 | -81.7949
2.1552 | -81.7937
2.1555 | -81.7907
2.1559 | -81.7858
2.1562 | -81.7843
2.1565 | -81.7834
2.1568 | -81.7814
2.1572 | -81.7802
2.1575 | -81.7788
2.1578 | -81.7755
2.1582 | -81.7735
2.1585 | -81.7715
2.1588 | -81.77
2.1591 | -81.7696
2.1595 | -81.7681
2.1598 | -81.7669
2.1601 | -81.765
2.1605 | -81.763
2.1608 | -81.7585
2.1611 | -81.7554
2.1614 | -81.7528
2.1618 | -81.7511
2.1621 | -81.7496
2.1624 | -81.7465
2.1628 | -81.7426
2.1631 | -81.7378
2.1634 | -81.7373
2.1638 | -81.7358
2.1641 | -81.7323
2.1644 | -81.7286
2.1647 | -81.7265
2.1651 | -81.7205
2.1654 | -81.7153
2.1657 | -81.7123
2.1661 | -81.7102
2.1664 | -81.7083
2.1667 | -81.7047
2.167 | -81.702
2.1674 | -81.6978
2.1677 | -81.6941
2.168 | -81.6899
2.1684 | -81.6848
2.1687 | -81.6776
2.169 | -81.6735
2.1693 | -81.6699
2.1697 | -81.6659
2.17 | -81.6598
2.1703 | -81.652
2.1707 | -81.646
2.171 | -81.6422
2.1713 | -81.6396
2.1717 | -81.6364
2.172 | -81.632
2.1723 | -81.6263
2.1726 | -81.622
2.173 | -81.6196
2.1733 | -81.6154
2.1736 | -81.6122
2.174 | -81.6113
2.1743 | -81.6072
2.1746 | -81.6008
2.1749 | -81.5971
2.1753 | -81.5931
2.1756 | -81.5871
2.1759 | -81.5798
2.1763 | -81.5746
2.1766 | -81.5677
2.1769 | -81.563
2.1772 | -81.5578
2.1776 | -81.5533
2.1779 | -81.5471
2.1782 | -81.5417
2.1786 | -81.5386
2.1789 | -81.5347
2.1792 | -81.529
2.1796 | -81.5245
2.1799 | -81.5185
2.1802 | -81.5129
2.1805 | -81.5078
2.1809 | -81.5028
2.1812 | -81.4985
2.1815 | -81.4947
2.1819 | -81.4905
2.1822 | -81.4871
2.1825 | -81.4835
2.1828 | -81.4784
2.1832 | -81.4721
2.1835 | -81.4667
2.1838 | -81.4619
2.1842 | -81.4557
2.1845 | -81.4534
2.1848 | -81.4486
2.1851 | -81.4433
2.1855 | -81.4383
2.1858 | -81.4328
2.1861 | -81.4271
2.1865 | -81.4217
2.1868 | -81.4169
2.1871 | -81.4124
2.1875 | -81.4079
2.1878 | -81.4026
2.1881 | -81.3981
2.1884 | -81.394
2.1888 | -81.3908
2.1891 | -81.3865
2.1894 | -81.382
2.1898 | -81.3791
2.1901 | -81.3743
2.1904 | -81.3688
2.1907 | -81.3652
2.1911 | -81.3616
2.1914 | -81.358
2.1917 | -81.3534
2.1921 | -81.3471
2.1924 | -81.3414
2.1927 | -81.3392
2.193 | -81.3338
2.1934 | -81.3286
2.1937 | -81.3231
2.194 | -81.3173
2.1944 | -81.3109
2.1947 | -81.3038
2.195 | -81.2989
2.1953 | -81.295
2.1957 | -81.2906
2.196 | -81.2876
2.1963 | -81.2849
2.1967 | -81.2805
2.197 | -81.2763
2.1973 | -81.271
2.1977 | -81.2668
2.198 | -81.2634
2.1983 | -81.2627
2.1986 | -81.2614
2.199 | -81.2615
2.1993 | -81.261
2.1996 | -81.2593
2.2 | -81.255
2.2003 | -81.2499
2.2006 | -81.2455
2.2009 | -81.2416
2.2013 | -81.2375
2.2016 | -81.2343
2.2019 | -81.2304
2.2023 | -81.2264
2.2026 | -81.2223
2.2029 | -81.2196
2.2032 | -81.2167
2.2036 | -81.2152
2.2039 | -81.2108
2.2042 | -81.2091
2.2046 | -81.2064
2.2049 | -81.2032
2.2052 | -81.1999
2.2056 | -81.1955
2.2059 | -81.1914
2.2062 | -81.1882
2.2065 | -81.1848
2.2069 | -81.1822
2.2072 | -81.1775
2.2075 | -81.1734
2.2079 | -81.1699
2.2082 | -81.1678
2.2085 | -81.1632
2.2088 | -81.16
2.2092 | -81.1573
2.2095 | -81.155
2.2098 | -81.1528
2.2102 | -81.151
2.2105 | -81.1495
2.2108 | -81.147
2.2111 | -81.1443
2.2115 | -81.1434
2.2118 | -81.1417
2.2121 | -81.1395
2.2125 | -81.1413
2.2128 | -81.1393
2.2131 | -81.1355
2.2135 | -81.133
2.2138 | -81.13
2.2141 | -81.1259
2.2144 | -81.1236
2.2148 | -81.1218
2.2151 | -81.1218
2.2154 | -81.1217
2.2158 | -81.1197
2.2161 | -81.1191
2.2164 | -81.1182
2.2167 | -81.1166
2.2171 | -81.1137
2.2174 | -81.1122
2.2177 | -81.1119
2.2181 | -81.1132
2.2184 | -81.1138
2.2187 | -81.1153
2.219 | -81.116
2.2194 | -81.116
2.2197 | -81.1175
2.22 | -81.1166
2.2204 | -81.1165
2.2207 | -81.1156
2.221 | -81.1138
2.2214 | -81.1137
2.2217 | -81.116
2.222 | -81.1202
2.2223 | -81.1218
2.2227 | -81.1242
2.223 | -81.1248
2.2233 | -81.1234
2.2237 | -81.1236
2.224 | -81.1237
2.2243 | -81.1226
2.2246 | -81.1208
2.225 | -81.1213
2.2253 | -81.1223
2.2256 | -81.123
2.226 | -81.1236
2.2263 | -81.1244
2.2266 | -81.1269
2.2269 | -81.1268
2.2273 | -81.1259
2.2276 | -81.1265
2.2279 | -81.1282
2.2283 | -81.1292
2.2286 | -81.1307
2.2289 | -81.133
2.2293 | -81.1356
2.2296 | -81.1365
2.2299 | -81.1365
2.2302 | -81.1381
2.2306 | -81.1424
2.2309 | -81.1447
2.2312 | -81.1459
2.2316 | -81.1522
2.2319 | -81.1545
2.2322 | -81.157
2.2325 | -81.1594
2.2329 | -81.162
2.2332 | -81.165
2.2335 | -81.1678
2.2339 | -81.1718
2.2342 | -81.1735
2.2345 | -81.1749
2.2348 | -81.1773
2.2352 | -81.1799
2.2355 | -81.1833
2.2358 | -81.1846
2.2362 | -81.1862
2.2365 | -81.1887
2.2368 | -81.1904
2.2372 | -81.1936
2.2375 | -81.1963
2.2378 | -81.1969
2.2381 | -81.1976
2.2385 | -81.1976
2.2388 | -81.1998
2.2391 | -81.2028
2.2395 | -81.2059
2.2398 | -81.2093
2.2401 | -81.2132
2.2404 | -81.2171
2.2408 | -81.2208
2.2411 | -81.2269
2.2414 | -81.2284
2.2418 | -81.2331
2.2421 | -81.2398
2.2424 | -81.2452
2.2427 | -81.2495
2.2431 | -81.2528
2.2434 | -81.2578
2.2437 | -81.2613
2.2441 | -81.2648
2.2444 | -81.267
2.2447 | -81.2677
2.245 | -81.2707
2.2454 | -81.2747
2.2457 | -81.2791
2.246 | -81.2829
2.2464 | -81.2824
2.2467 | -81.2816
2.247 | -81.2805
2.2474 | -81.2807
2.2477 | -81.2838
2.248 | -81.2874
2.2483 | -81.2914
2.2487 | -81.2961
2.249 | -81.2988
2.2493 | -81.3031
2.2497 | -81.308
2.25 | -81.3125
2.2503 | -81.3153
2.2506 | -81.315
2.251 | -81.3208
2.2513 | -81.3266
2.2516 | -81.3309
2.252 | -81.3358
2.2523 | -81.34
2.2526 | -81.343
2.2529 | -81.3444
2.2533 | -81.3461
2.2536 | -81.3507
2.2539 | -81.3541
2.2543 | -81.3584
2.2546 | -81.3635
2.2549 | -81.3681
2.2553 | -81.3711
2.2556 | -81.3741
2.2559 | -81.3771
2.2562 | -81.3796
2.2566 | -81.3832
2.2569 | -81.3867
2.2572 | -81.3921
2.2576 | -81.3971
2.2579 | -81.4011
2.2582 | -81.4048
2.2585 | -81.4074
2.2589 | -81.4094
2.2592 | -81.4129
2.2595 | -81.4145
2.2599 | -81.414
2.2602 | -81.4162
2.2605 | -81.4194
2.2608 | -81.4232
2.2612 | -81.4236
2.2615 | -81.4259
2.2618 | -81.4299
2.2622 | -81.4327
2.2625 | -81.4366
2.2628 | -81.4387
2.2632 | -81.4406
2.2635 | -81.443
2.2638 | -81.4483
2.2641 | -81.4548
2.2645 | -81.4586
2.2648 | -81.462
2.2651 | -81.4644
2.2655 | -81.4668
2.2658 | -81.4692
2.2661 | -81.4724
2.2664 | -81.4761
2.2668 | -81.4793
2.2671 | -81.482
2.2674 | -81.4833
2.2678 | -81.4861
2.2681 | -81.4889
2.2684 | -81.4928
2.2687 | -81.4964
2.2691 | -81.4994
2.2694 | -81.5014
2.2697 | -81.5031
2.2701 | -81.507
2.2704 | -81.5102
2.2707 | -81.5132
2.2711 | -81.5141
2.2714 | -81.5161
2.2717 | -81.5198
2.272 | -81.5211
2.2724 | -81.5203
2.2727 | -81.5206
2.273 | -81.5218
2.2734 | -81.5211
2.2737 | -81.523
2.274 | -81.5244
2.2743 | -81.5238
2.2747 | -81.5254
2.275 | -81.5327
2.2753 | -81.5335
2.2757 | -81.5344
2.276 | -81.5375
2.2763 | -81.5381
2.2766 | -81.5381
2.277 | -81.5388
2.2773 | -81.5391
2.2776 | -81.5402
2.278 | -81.5411
2.2783 | -81.5414
2.2786 | -81.5434
2.279 | -81.5421
2.2793 | -81.5452
2.2796 | -81.5454
2.2799 | -81.5462
2.2803 | -81.5464
2.2806 | -81.5517
2.2809 | -81.5605
2.2813 | -81.5626
2.2816 | -81.5624
2.2819 | -81.5619
2.2822 | -81.56
2.2826 | -81.5602
2.2829 | -81.5604
2.2832 | -81.56
2.2836 | -81.5605
2.2839 | -81.5602
2.2842 | -81.5591
2.2845 | -81.5579
2.2849 | -81.5564
2.2852 | -81.5546
2.2855 | -81.5546
2.2859 | -81.554
2.2862 | -81.5529
2.2865 | -81.552
2.2868 | -81.5514
2.2872 | -81.5509
2.2875 | -81.5508
2.2878 | -81.5524
2.2882 | -81.5464
2.2885 | -81.5445
2.2888 | -81.5425
2.2892 | -81.5398
2.2895 | -81.5385
2.2898 | -81.5382
2.2901 | -81.5366
2.2905 | -81.5356
2.2908 | -81.5361
2.2911 | -81.5348
2.2915 | -81.5332
2.2918 | -81.5307
2.2921 | -81.529
2.2924 | -81.5274
2.2928 | -81.5264
2.2931 | -81.5271
2.2934 | -81.5254
2.2938 | -81.5242
2.2941 | -81.5206
2.2944 | -81.5187
2.2948 | -81.517
2.2951 | -81.5155
2.2954 | -81.5148
2.2957 | -81.5114
2.2961 | -81.5088
2.2964 | -81.5041
2.2967 | -81.5
2.2971 | -81.4988
2.2974 | -81.498
2.2977 | -81.4974
2.298 | -81.4958
2.2984 | -81.4941
2.2987 | -81.4921
2.299 | -81.4897
2.2994 | -81.4875
2.2997 | -81.4884
2.3 | -81.4865
2.3003 | -81.4843
2.3007 | -81.4844
2.301 | -81.4867
2.3013 | -81.4841
2.3017 | -81.4807
2.302 | -81.4788
2.3023 | -81.4775
2.3026 | -81.4753
2.303 | -81.4719
2.3033 | -81.4697
2.3036 | -81.4681
2.304 | -81.4663
2.3043 | -81.4659
2.3046 | -81.4625
2.305 | -81.4583
2.3053 | -81.4556
2.3056 | -81.4519
2.3059 | -81.4483
2.3063 | -81.446
2.3066 | -81.442
2.3069 | -81.4373
2.3073 | -81.4344
2.3076 | -81.4316
2.3079 | -81.4295
2.3082 | -81.4272
2.3086 | -81.4242
2.3089 | -81.4218
2.3092 | -81.4186
2.3096 | -81.4157
2.3099 | -81.4127
2.3102 | -81.4098
2.3105 | -81.4071
2.3109 | -81.4042
2.3112 | -81.401
2.3115 | -81.4017
2.3119 | -81.3993
2.3122 | -81.3972
2.3125 | -81.3921
2.3129 | -81.3887
2.3132 | -81.3866
2.3135 | -81.3864
2.3138 | -81.3851
2.3142 | -81.3818
2.3145 | -81.3785
2.3148 | -81.3756
2.3152 | -81.3732
2.3155 | -81.3704
2.3158 | -81.3669
2.3161 | -81.3621
2.3165 | -81.3623
2.3168 | -81.3621
2.3171 | -81.3601
2.3175 | -81.358
2.3178 | -81.3558
2.3181 | -81.3533
2.3184 | -81.3503
2.3188 | -81.3502
2.3191 | -81.35
2.3194 | -81.3471
2.3198 | -81.3439
2.3201 | -81.3427
2.3204 | -81.342
2.3208 | -81.3401
2.3211 | -81.3385
2.3214 | -81.338
2.3217 | -81.3371
2.3221 | -81.3353
2.3224 | -81.333
2.3227 | -81.3319
2.3231 | -81.3323
2.3234 | -81.3315
2.3237 | -81.3305
2.324 | -81.3291
2.3244 | -81.3295
2.3247 | -81.3307
2.325 | -81.333
2.3254 | -81.3317
2.3257 | -81.329
2.326 | -81.3269
2.3263 | -81.3269
2.3267 | -81.3269
2.327 | -81.3287
2.3273 | -81.3287
2.3277 | -81.3267
2.328 | -81.3244
2.3283 | -81.3225
2.3287 | -81.3196
2.329 | -81.3172
2.3293 | -81.3167
2.3296 | -81.3151
2.33 | -81.3129
2.3303 | -81.3117
2.3306 | -81.3112
2.331 | -81.3098
2.3313 | -81.3097
2.3316 | -81.3081
2.3319 | -81.3063
2.3323 | -81.3058
2.3326 | -81.3066
2.3329 | -81.3073
2.3333 | -81.3075
2.3336 | -81.3085
2.3339 | -81.308
2.3342 | -81.3073
2.3346 | -81.3069
2.3349 | -81.3082
2.3352 | -81.3083
2.3356 | -81.3075
2.3359 | -81.3086
2.3362 | -81.3057
2.3365 | -81.3058
2.3369 | -81.3064
2.3372 | -81.3084
2.3375 | -81.3082
2.3379 | -81.3062
2.3382 | -81.3059
2.3385 | -81.3049
2.3389 | -81.3046
2.3392 | -81.3054
2.3395 | -81.306
2.3398 | -81.3058
2.3402 | -81.3051
2.3405 | -81.3064
2.3408 | -81.3071
2.3412 | -81.3073
2.3415 | -81.3086
2.3418 | -81.3101
2.3421 | -81.3113
2.3425 | -81.3114
2.3428 | -81.3115
2.3431 | -81.3123
2.3435 | -81.313
2.3438 | -81.3144
2.3441 | -81.3157
2.3444 | -81.3171
2.3448 | -81.3159
2.3451 | -81.3151
2.3454 | -81.3188
2.3458 | -81.32
2.3461 | -81.3196
2.3464 | -81.3185
2.3468 | -81.318
2.3471 | -81.3198
2.3474 | -81.3205
2.3477 | -81.3189
2.3481 | -81.3201
2.3484 | -81.3217
2.3487 | -81.3229
2.3491 | -81.324
2.3494 | -81.3249
2.3497 | -81.3272
2.35 | -81.3299
2.3504 | -81.3336
2.3507 | -81.3359
2.351 | -81.3378
2.3514 | -81.3397
2.3517 | -81.3414
2.352 | -81.3359
2.3523 | -81.3249
2.3527 | -81.3225
2.353 | -81.326
2.3533 | -81.3276
2.3537 | -81.3272
2.354 | -81.3268
2.3543 | -81.3286
2.3547 | -81.3295
2.355 | -81.3331
2.3553 | -81.3345
2.3556 | -81.3336
2.356 | -81.3353
2.3563 | -81.3367
2.3566 | -81.3373
2.357 | -81.338
2.3573 | -81.3395
2.3576 | -81.3414
2.3579 | -81.3421
2.3583 | -81.3412
2.3586 | -81.3412
2.3589 | -81.3409
2.3593 | -81.3412
2.3596 | -81.3414
2.3599 | -81.3411
2.3602 | -81.3417
2.3606 | -81.3435
2.3609 | -81.345
2.3612 | -81.347
2.3616 | -81.3499
2.3619 | -81.3521
2.3622 | -81.3538
2.3626 | -81.3541
2.3629 | -81.3556
2.3632 | -81.3569
2.3635 | -81.3585
2.3639 | -81.3594
2.3642 | -81.3607
2.3645 | -81.3614
2.3649 | -81.3607
2.3652 | -81.3598
2.3655 | -81.3606
2.3658 | -81.3632
2.3662 | -81.363
2.3665 | -81.3649
2.3668 | -81.3679
2.3672 | -81.3699
2.3675 | -81.3674
2.3678 | -81.3698
2.3681 | -81.3715
2.3685 | -81.3727
2.3688 | -81.3727
2.3691 | -81.3727
2.3695 | -81.3723
2.3698 | -81.3723
2.3701 | -81.3736
2.3705 | -81.3747
2.3708 | -81.3748
2.3711 | -81.3754
2.3714 | -81.3761
2.3718 | -81.3773
2.3721 | -81.3778
2.3724 | -81.3791
2.3728 | -81.3816
2.3731 | -81.3826
2.3734 | -81.3848
2.3737 | -81.385
2.3741 | -81.3834
2.3744 | -81.3844
2.3747 | -81.3831
2.3751 | -81.3835
2.3754 | -81.3814
2.3757 | -81.3816
2.376 | -81.3821
2.3764 | -81.3822
2.3767 | -81.383
2.377 | -81.3823
2.3774 | -81.3805
2.3777 | -81.3797
2.378 | -81.3789
2.3784 | -81.3791
2.3787 | -81.3792
2.379 | -81.3789
2.3793 | -81.3788
2.3797 | -81.3794
2.38 | -81.381
2.3803 | -81.3819
2.3807 | -81.3834
2.381 | -81.3844
2.3813 | -81.3844
2.3816 | -81.3841
2.382 | -81.3839
2.3823 | -81.3836
2.3826 | -81.384
2.383 | -81.3844
2.3833 | -81.3837
2.3836 | -81.3831
2.3839 | -81.3813
2.3843 | -81.3797
2.3846 | -81.3772
2.3849 | -81.3786
2.3853 | -81.3792
2.3856 | -81.3789
2.3859 | -81.3776
2.3862 | -81.3767
2.3866 | -81.3741
2.3869 | -81.372
2.3872 | -81.3711
2.3876 | -81.3718
2.3879 | -81.3716
2.3882 | -81.3713
2.3886 | -81.3698
2.3889 | -81.3692
2.3892 | -81.367
2.3895 | -81.3628
2.3899 | -81.3582
2.3902 | -81.3554
2.3905 | -81.3561
2.3909 | -81.3559
2.3912 | -81.3526
2.3915 | -81.3517
2.3918 | -81.3508
2.3922 | -81.3487
2.3925 | -81.344
2.3928 | -81.3391
2.3932 | -81.3343
2.3935 | -81.3318
2.3938 | -81.3304
2.3941 | -81.3279
2.3945 | -81.3258
2.3948 | -81.3235
2.3951 | -81.3207
2.3955 | -81.3211
2.3958 | -81.3205
2.3961 | -81.3183
2.3965 | -81.3165
2.3968 | -81.3149
2.3971 | -81.3146
2.3974 | -81.3139
2.3978 | -81.3128
2.3981 | -81.312
2.3984 | -81.3108
2.3988 | -81.3099
2.3991 | -81.3099
2.3994 | -81.3084
2.3997 | -81.3066
2.4001 | -81.3048
2.4004 | -81.3023
2.4007 | -81.3004
2.4011 | -81.2979
2.4014 | -81.295
2.4017 | -81.2917
2.402 | -81.2896
2.4024 | -81.2868
2.4027 | -81.2839
2.403 | -81.2811
2.4034 | -81.2783
2.4037 | -81.2759
2.404 | -81.2724
2.4044 | -81.27
2.4047 | -81.2656
2.405 | -81.2635
2.4053 | -81.2609
2.4057 | -81.2588
2.406 | -81.2557
2.4063 | -81.2523
2.4067 | -81.252
2.407 | -81.2478
2.4073 | -81.2443
2.4076 | -81.2418
2.408 | -81.2387
2.4083 | -81.2349
2.4086 | -81.2318
2.409 | -81.2336
2.4093 | -81.2334
2.4096 | -81.2308
2.4099 | -81.2273
2.4103 | -81.2236
2.4106 | -81.2208
2.4109 | -81.2192
2.4113 | -81.2161
2.4116 | -81.2134
2.4119 | -81.2098
2.4123 | -81.2052
2.4126 | -81.2014
2.4129 | -81.1998
2.4132 | -81.198
2.4136 | -81.1937
2.4139 | -81.1907
2.4142 | -81.1876
2.4146 | -81.1845
2.4149 | -81.1818
2.4152 | -81.178
2.4155 | -81.1742
2.4159 | -81.1719
2.4162 | -81.1689
2.4165 | -81.1645
2.4169 | -81.1603
2.4172 | -81.1569
2.4175 | -81.1528
2.4178 | -81.1489
2.4182 | -81.1457
2.4185 | -81.1432
2.4188 | -81.1391
2.4192 | -81.1293
2.4195 | -81.1216
2.4198 | -81.1186
2.4202 | -81.1148
2.4205 | -81.1103
2.4208 | -81.1062
2.4211 | -81.1038
2.4215 | -81.1018
2.4218 | -81.1008
2.4221 | -81.0972
2.4225 | -81.0954
2.4228 | -81.0931
2.4231 | -81.091
2.4234 | -81.0882
2.4238 | -81.085
2.4241 | -81.0813
2.4244 | -81.0777
2.4248 | -81.0741
2.4251 | -81.0706
2.4254 | -81.0679
2.4257 | -81.065
2.4261 | -81.0631
2.4264 | -81.0599
2.4267 | -81.0567
2.4271 | -81.052
2.4274 | -81.0457
2.4277 | -81.0405
2.428 | -81.0362
2.4284 | -81.0327
2.4287 | -81.0273
2.429 | -81.0224
2.4294 | -81.0161
2.4297 | -81.0123
2.43 | -81.0092
2.4304 | -81.0059
2.4307 | -81.0032
2.431 | -81.0024
2.4313 | -80.9987
2.4317 | -80.9901
2.432 | -80.9852
2.4323 | -80.9818
2.4327 | -80.9782
2.433 | -80.9753
2.4333 | -80.9721
2.4336 | -80.9689
2.434 | -80.9665
2.4343 | -80.9637
2.4346 | -80.9577
2.435 | -80.9529
2.4353 | -80.9501
2.4356 | -80.9471
2.436 | -80.9464
2.4363 | -80.9464
2.4366 | -80.9437
2.4369 | -80.9414
2.4373 | -80.9406
2.4376 | -80.938
2.4379 | -80.9354
2.4383 | -80.9343
2.4386 | -80.9326
2.4389 | -80.9301
2.4392 | -80.93
2.4396 | -80.9276
2.4399 | -80.9251
2.4402 | -80.9219
2.4406 | -80.9219
2.4409 | -80.9202
2.4412 | -80.9171
2.4415 | -80.9141
2.4419 | -80.9117
2.4422 | -80.9093
2.4425 | -80.9077
2.4429 | -80.9052
2.4432 | -80.9028
2.4435 | -80.9011
2.4438 | -80.8989
2.4442 | -80.8969
2.4445 | -80.8932
2.4448 | -80.8912
2.4452 | -80.8884
2.4455 | -80.8853
2.4458 | -80.8818
2.4462 | -80.8796
2.4465 | -80.8775
2.4468 | -80.873
2.4471 | -80.8701
2.4475 | -80.8675
2.4478 | -80.864
2.4481 | -80.8612
2.4485 | -80.8589
2.4488 | -80.858
2.4491 | -80.8562
2.4494 | -80.8539
2.4498 | -80.8515
2.4501 | -80.8503
2.4504 | -80.8465
2.4508 | -80.8443
2.4511 | -80.8426
2.4514 | -80.841
2.4517 | -80.8394
2.4521 | -80.8378
2.4524 | -80.8372
2.4527 | -80.8357
2.4531 | -80.8377
2.4534 | -80.8363
2.4537 | -80.8348
2.4541 | -80.8332
2.4544 | -80.8314
2.4547 | -80.8297
2.455 | -80.8279
2.4554 | -80.8274
2.4557 | -80.8258
2.456 | -80.8237
2.4564 | -80.821
2.4567 | -80.8205
2.457 | -80.8207
2.4573 | -80.8194
2.4577 | -80.8182
2.458 | -80.8174
2.4583 | -80.8166
2.4587 | -80.816
2.459 | -80.8122
2.4593 | -80.8094
2.4596 | -80.8092
2.46 | -80.807
2.4603 | -80.8063
2.4606 | -80.803
2.461 | -80.8016
2.4613 | -80.8021
2.4616 | -80.8006
2.462 | -80.8014
2.4623 | -80.8017
2.4626 | -80.8009
2.4629 | -80.7984
2.4633 | -80.7977
2.4636 | -80.7923
2.4639 | -80.7893
2.4643 | -80.7879
2.4646 | -80.7856
2.4649 | -80.7852
2.4652 | -80.7868
2.4656 | -80.7855
2.4659 | -80.7839
2.4662 | -80.7826
2.4666 | -80.7817
2.4669 | -80.7806
2.4672 | -80.7801
2.4675 | -80.7799
2.4679 | -80.7791
2.4682 | -80.7785
2.4685 | -80.7776
2.4689 | -80.7785
2.4692 | -80.7788
2.4695 | -80.7787
2.4699 | -80.7794
2.4702 | -80.7817
2.4705 | -80.7826
2.4708 | -80.7815
2.4712 | -80.7807
2.4715 | -80.7798
2.4718 | -80.7793
2.4722 | -80.779
2.4725 | -80.7784
2.4728 | -80.7794
2.4731 | -80.7795
2.4735 | -80.7785
2.4738 | -80.7772
2.4741 | -80.7762
2.4745 | -80.7747
2.4748 | -80.7736
2.4751 | -80.7731
2.4754 | -80.7735
2.4758 | -80.7735
2.4761 | -80.773
2.4764 | -80.7733
2.4768 | -80.7732
2.4771 | -80.7724
2.4774 | -80.7724
2.4777 | -80.7727
2.4781 | -80.7716
2.4784 | -80.7725
2.4787 | -80.7707
2.4791 | -80.7705
2.4794 | -80.7618
2.4797 | -80.7541
2.4801 | -80.7523
2.4804 | -80.7501
2.4807 | -80.7474
2.481 | -80.7475
2.4814 | -80.7468
2.4817 | -80.7463
2.482 | -80.7456
2.4824 | -80.7451
2.4827 | -80.7441
2.483 | -80.7437
2.4833 | -80.7428
2.4837 | -80.741
2.484 | -80.7408
2.4843 | -80.7411
2.4847 | -80.7402
2.485 | -80.7385
2.4853 | -80.7383
2.4857 | -80.7381
2.486 | -80.7366
2.4863 | -80.7356
2.4866 | -80.7361
2.487 | -80.7367
2.4873 | -80.7361
2.4876 | -80.7356
2.488 | -80.7363
2.4883 | -80.7358
2.4886 | -80.7331
2.4889 | -80.7326
2.4893 | -80.7311
2.4896 | -80.7295
2.4899 | -80.7297
2.4903 | -80.7282
2.4906 | -80.7288
2.4909 | -80.7285
2.4912 | -80.7282
2.4916 | -80.7272
2.4919 | -80.726
2.4922 | -80.7256
2.4926 | -80.7263
2.4929 | -80.7274
2.4932 | -80.7277
2.4935 | -80.7274
2.4939 | -80.7273
2.4942 | -80.728
2.4945 | -80.7275
2.4949 | -80.7276
2.4952 | -80.7275
2.4955 | -80.7267
2.4959 | -80.726
2.4962 | -80.7256
2.4965 | -80.724
2.4968 | -80.7257
2.4972 | -80.7245
2.4975 | -80.7249
2.4978 | -80.7246
2.4982 | -80.7247
2.4985 | -80.7245
2.4988 | -80.7236
2.4991 | -80.7238
2.4995 | -80.7208
2.4998 | -80.7196
2.5001 | -80.719
2.5005 | -80.7179
2.5008 | -80.7173
2.5011 | -80.7163
2.5014 | -80.7156
2.5018 | -80.7161
2.5021 | -80.7161
2.5024 | -80.7164
2.5028 | -80.7148
2.5031 | -80.7131
2.5034 | -80.7106
2.5038 | -80.7091
2.5041 | -80.7093
2.5044 | -80.7089
2.5047 | -80.7079
2.5051 | -80.7075
2.5054 | -80.7068
2.5057 | -80.7059
2.5061 | -80.7066
2.5064 | -80.7078
2.5067 | -80.7067
2.507 | -80.7077
2.5074 | -80.7077
2.5077 | -80.7058
2.508 | -80.7047
2.5084 | -80.7042
2.5087 | -80.7039
2.509 | -80.7056
2.5093 | -80.7057
2.5097 | -80.7052
2.51 | -80.7054
2.5103 | -80.705
2.5107 | -80.7018
2.511 | -80.7019
2.5113 | -80.7019
2.5117 | -80.7019
2.512 | -80.7012
2.5123 | -80.7004
2.5126 | -80.6987
2.513 | -80.6984
2.5133 | -80.6987
2.5136 | -80.7009
2.514 | -80.7023
2.5143 | -80.7016
2.5146 | -80.7005
2.5149 | -80.7003
2.5153 | -80.7013
2.5156 | -80.7031
2.5159 | -80.7046
2.5163 | -80.7016
2.5166 | -80.6999
2.5169 | -80.6993
2.5172 | -80.7
2.5176 | -80.698
2.5179 | -80.6967
2.5182 | -80.6966
2.5186 | -80.697
2.5189 | -80.6973
2.5192 | -80.6966
2.5196 | -80.6965
2.5199 | -80.6953
2.5202 | -80.6933
2.5205 | -80.6925
2.5209 | -80.6905
2.5212 | -80.6895
2.5215 | -80.6958
2.5219 | -80.7058
2.5222 | -80.7089
2.5225 | -80.7073
2.5228 | -80.706
2.5232 | -80.7059
2.5235 | -80.706
2.5238 | -80.7046
2.5242 | -80.7039
2.5245 | -80.7028
2.5248 | -80.701
2.5251 | -80.6989
2.5255 | -80.6984
2.5258 | -80.698
2.5261 | -80.6969
2.5265 | -80.6955
2.5268 | -80.6942
2.5271 | -80.6943
2.5274 | -80.6918
2.5278 | -80.6886
2.5281 | -80.6875
2.5284 | -80.6883
2.5288 | -80.6883
2.5291 | -80.69
2.5294 | -80.6905
2.5298 | -80.6897
2.5301 | -80.6885
2.5304 | -80.688
2.5307 | -80.687
2.5311 | -80.686
2.5314 | -80.6831
2.5317 | -80.6808
2.5321 | -80.6772
2.5324 | -80.6745
2.5327 | -80.6741
2.533 | -80.6719
2.5334 | -80.6682
2.5337 | -80.6667
2.534 | -80.6645
2.5344 | -80.6638
2.5347 | -80.6628
2.535 | -80.6621
2.5353 | -80.6613
2.5357 | -80.6629
2.536 | -80.664
2.5363 | -80.6625
2.5367 | -80.6614
2.537 | -80.6606
2.5373 | -80.659
2.5377 | -80.6574
2.538 | -80.6567
2.5383 | -80.6584
2.5386 | -80.6569
2.539 | -80.6518
2.5393 | -80.6463
2.5396 | -80.6429
2.54 | -80.6416
2.5403 | -80.6415
2.5406 | -80.6402
2.5409 | -80.6377
2.5413 | -80.6362
2.5416 | -80.6356
2.5419 | -80.6344
2.5423 | -80.633
2.5426 | -80.6317
2.5429 | -80.6288
2.5432 | -80.6271
2.5436 | -80.6231
2.5439 | -80.6181
2.5442 | -80.6145
2.5446 | -80.6145
2.5449 | -80.612
2.5452 | -80.6112
2.5456 | -80.6097
2.5459 | -80.6076
2.5462 | -80.6076
2.5465 | -80.6063
2.5469 | -80.6067
2.5472 | -80.6043
2.5475 | -80.6027
2.5479 | -80.5994
2.5482 | -80.5975
2.5485 | -80.596
2.5488 | -80.596
2.5492 | -80.5938
2.5495 | -80.5907
2.5498 | -80.588
2.5502 | -80.5861
2.5505 | -80.5846
2.5508 | -80.58
2.5511 | -80.5728
2.5515 | -80.5705
2.5518 | -80.5681
2.5521 | -80.566
2.5525 | -80.5645
2.5528 | -80.5622
2.5531 | -80.5604
2.5535 | -80.5583
2.5538 | -80.557
2.5541 | -80.5553
2.5544 | -80.5533
2.5548 | -80.5502
2.5551 | -80.5482
2.5554 | -80.5437
2.5558 | -80.5393
2.5561 | -80.5372
2.5564 | -80.5344
2.5567 | -80.5346
2.5571 | -80.5329
2.5574 | -80.5313
2.5577 | -80.5278
2.5581 | -80.5264
2.5584 | -80.5232
2.5587 | -80.5222
2.559 | -80.5208
2.5594 | -80.5184
2.5597 | -80.5157
2.56 | -80.5127
2.5604 | -80.5096
2.5607 | -80.5077
2.561 | -80.5048
2.5614 | -80.502
2.5617 | -80.5024
2.562 | -80.5029
2.5623 | -80.5022
2.5627 | -80.5002
2.563 | -80.4978
2.5633 | -80.4963
2.5637 | -80.494
2.564 | -80.4927
2.5643 | -80.4895
2.5646 | -80.4873
2.565 | -80.4866
2.5653 | -80.4835
2.5656 | -80.4815
2.566 | -80.4817
2.5663 | -80.4802
2.5666 | -80.4775
2.5669 | -80.4751
2.5673 | -80.4733
2.5676 | -80.4721
2.5679 | -80.4687
2.5683 | -80.4659
2.5686 | -80.4641
2.5689 | -80.4598
2.5693 | -80.456
2.5696 | -80.4539
2.5699 | -80.4522
2.5702 | -80.45
2.5706 | -80.448
2.5709 | -80.4463
2.5712 | -80.4442
2.5716 | -80.4424
2.5719 | -80.4402
2.5722 | -80.4397
2.5725 | -80.4397
2.5729 | -80.4375
2.5732 | -80.4338
2.5735 | -80.4275
2.5739 | -80.4241
2.5742 | -80.422
2.5745 | -80.4206
2.5748 | -80.4195
2.5752 | -80.4167
2.5755 | -80.4136
2.5758 | -80.4117
2.5762 | -80.4087
2.5765 | -80.4054
2.5768 | -80.4046
2.5772 | -80.4033
2.5775 | -80.402
2.5778 | -80.4001
2.5781 | -80.3997
2.5785 | -80.3982
2.5788 | -80.3961
2.5791 | -80.3942
2.5795 | -80.3912
2.5798 | -80.3894
2.5801 | -80.3848
2.5804 | -80.3844
2.5808 | -80.3831
2.5811 | -80.3805
2.5814 | -80.3794
2.5818 | -80.3814
2.5821 | -80.3808
2.5824 | -80.3786
2.5827 | -80.3785
2.5831 | -80.3766
2.5834 | -80.3749
2.5837 | -80.3732
2.5841 | -80.3714
2.5844 | -80.3692
2.5847 | -80.3669
2.585 | -80.3648
2.5854 | -80.363
2.5857 | -80.3618
2.586 | -80.3594
2.5864 | -80.3566
2.5867 | -80.3539
2.587 | -80.3529
2.5874 | -80.3499
2.5877 | -80.3493
2.588 | -80.3461
2.5883 | -80.3434
2.5887 | -80.3415
2.589 | -80.3396
2.5893 | -80.3394
2.5897 | -80.3392
2.59 | -80.3382
2.5903 | -80.3366
2.5906 | -80.3364
2.591 | -80.3342
2.5913 | -80.3324
2.5916 | -80.331
2.592 | -80.33
2.5923 | -80.3306
2.5926 | -80.3291
2.5929 | -80.3282
2.5933 | -80.3287
2.5936 | -80.3319
2.5939 | -80.3348
2.5943 | -80.3326
2.5946 | -80.3312
2.5949 | -80.3279
2.5953 | -80.3236
2.5956 | -80.3206
2.5959 | -80.3196
2.5962 | -80.3166
2.5966 | -80.3154
2.5969 | -80.3132
2.5972 | -80.3114
2.5976 | -80.3111
2.5979 | -80.3104
2.5982 | -80.3093
2.5985 | -80.3083
2.5989 | -80.3071
2.5992 | -80.3081
2.5995 | -80.3072
2.5999 | -80.3058
2.6002 | -80.304
2.6005 | -80.3009
2.6008 | -80.3002
2.6012 | -80.3003
2.6015 | -80.3004
2.6018 | -80.3014
2.6022 | -80.3025
2.6025 | -80.3037
2.6028 | -80.303
2.6032 | -80.3016
2.6035 | -80.2996
2.6038 | -80.2983
2.6041 | -80.2964
2.6045 | -80.2972
2.6048 | -80.2964
2.6051 | -80.2958
2.6055 | -80.2957
2.6058 | -80.2945
2.6061 | -80.2917
2.6064 | -80.2902
2.6068 | -80.2899
2.6071 | -80.2931
2.6074 | -80.2936
2.6078 | -80.2928
2.6081 | -80.2927
2.6084 | -80.2919
2.6087 | -80.289
2.6091 | -80.2849
2.6094 | -80.2821
2.6097 | -80.2797
2.6101 | -80.2791
2.6104 | -80.2807
2.6107 | -80.2755
2.6111 | -80.2693
2.6114 | -80.2631
2.6117 | -80.2589
2.612 | -80.2574
2.6124 | -80.2564
2.6127 | -80.2538
2.613 | -80.2529
2.6134 | -80.251
2.6137 | -80.2492
2.614 | -80.2482
2.6143 | -80.2474
2.6147 | -80.2467
2.615 | -80.2451
2.6153 | -80.2438
2.6157 | -80.2417
2.616 | -80.2421
2.6163 | -80.2399
2.6166 | -80.2385
2.617 | -80.2368
2.6173 | -80.2357
2.6176 | -80.2323
2.618 | -80.2309
2.6183 | -80.2293
2.6186 | -80.2282
2.6189 | -80.2277
2.6193 | -80.2265
2.6196 | -80.225
2.6199 | -80.2245
2.6203 | -80.2237
2.6206 | -80.2245
2.6209 | -80.2252
2.6213 | -80.2246
2.6216 | -80.2244
2.6219 | -80.223
2.6222 | -80.2206
2.6226 | -80.2196
2.6229 | -80.2181
2.6232 | -80.219
2.6236 | -80.2179
2.6239 | -80.2171
2.6242 | -80.2171
2.6245 | -80.2158
2.6249 | -80.2132
2.6252 | -80.2101
2.6255 | -80.2084
2.6259 | -80.208
2.6262 | -80.2077
2.6265 | -80.2055
2.6269 | -80.2032
2.6272 | -80.1985
2.6275 | -80.1957
2.6278 | -80.1938
2.6282 | -80.1901
2.6285 | -80.1864
2.6288 | -80.1839
2.6292 | -80.1815
2.6295 | -80.1798
2.6298 | -80.1789
2.6301 | -80.1769
2.6305 | -80.1748
2.6308 | -80.1728
2.6311 | -80.1708
2.6315 | -80.1695
2.6318 | -80.1678
2.6321 | -80.1661
2.6324 | -80.1649
2.6328 | -80.1629
2.6331 | -80.1609
2.6334 | -80.1615
2.6338 | -80.1576
2.6341 | -80.1558
2.6344 | -80.1548
2.6347 | -80.1521
2.6351 | -80.1483
2.6354 | -80.1467
2.6357 | -80.146
2.6361 | -80.1468
2.6364 | -80.147
2.6367 | -80.1456
2.6371 | -80.1447
2.6374 | -80.1442
2.6377 | -80.1416
2.638 | -80.1378
2.6384 | -80.1345
2.6387 | -80.1326
2.639 | -80.1317
2.6394 | -80.1306
2.6397 | -80.1287
2.64 | -80.1261
2.6403 | -80.123
2.6407 | -80.1192
2.641 | -80.1155
2.6413 | -80.1167
2.6417 | -80.1165
2.642 | -80.1145
2.6423 | -80.114
2.6426 | -80.1126
2.643 | -80.1093
2.6433 | -80.1075
2.6436 | -80.1054
2.644 | -80.103
2.6443 | -80.0981
2.6446 | -80.0941
2.645 | -80.0895
2.6453 | -80.0863
2.6456 | -80.084
2.6459 | -80.0817
2.6463 | -80.08
2.6466 | -80.0762
2.6469 | -80.0731
2.6473 | -80.071
2.6476 | -80.0685
2.6479 | -80.0658
2.6482 | -80.0605
2.6486 | -80.0566
2.6489 | -80.053
2.6492 | -80.0482
2.6496 | -80.0446
2.6499 | -80.0397
2.6502 | -80.0366
2.6505 | -80.0328
2.6509 | -80.028
2.6512 | -80.0234
2.6515 | -80.0201
2.6519 | -80.0173
2.6522 | -80.0153
2.6525 | -80.0127
2.6529 | -80.0103
2.6532 | -80.0071
2.6535 | -80.0039
2.6538 | -79.9999
2.6542 | -79.9956
2.6545 | -79.9907
2.6548 | -79.9869
2.6552 | -79.9834
2.6555 | -79.9796
2.6558 | -79.976
2.6561 | -79.9733
2.6565 | -79.9657
2.6568 | -79.9596
2.6571 | -79.9568
2.6575 | -79.9548
2.6578 | -79.9508
2.6581 | -79.9531
2.6584 | -79.953
2.6588 | -79.9458
2.6591 | -79.9388
2.6594 | -79.9344
2.6598 | -79.9254
2.6601 | -79.916
2.6604 | -79.9116
2.6608 | -79.9068
2.6611 | -79.9005
2.6614 | -79.8932
2.6617 | -79.8897
2.6621 | -79.8854
2.6624 | -79.8811
2.6627 | -79.8767
2.6631 | -79.8719
2.6634 | -79.8637
2.6637 | -79.8544
2.664 | -79.8452
2.6644 | -79.8402
2.6647 | -79.8346
2.665 | -79.83
2.6654 | -79.8247
2.6657 | -79.8185
2.666 | -79.8136
2.6663 | -79.8067
2.6667 | -79.7983
2.667 | -79.7927
2.6673 | -79.7867
2.6677 | -79.781
2.668 | -79.7762
2.6683 | -79.7698
2.6686 | -79.7648
2.669 | -79.7605
2.6693 | -79.7556
2.6696 | -79.748
2.67 | -79.7418
2.6703 | -79.7355
2.6706 | -79.7295
2.671 | -79.7242
2.6713 | -79.7193
2.6716 | -79.714
2.6719 | -79.706
2.6723 | -79.7005
2.6726 | -79.6935
2.6729 | -79.6873
2.6733 | -79.6808
2.6736 | -79.6749
2.6739 | -79.6682
2.6742 | -79.6621
2.6746 | -79.6562
2.6749 | -79.65
2.6752 | -79.6432
2.6756 | -79.6369
2.6759 | -79.631
2.6762 | -79.6266
2.6766 | -79.6206
2.6769 | -79.615
2.6772 | -79.6096
2.6775 | -79.6032
2.6779 | -79.5965
2.6782 | -79.5909
2.6785 | -79.5854
2.6789 | -79.58
2.6792 | -79.5735
2.6795 | -79.5676
2.6798 | -79.5637
2.6802 | -79.5587
2.6805 | -79.5521
2.6808 | -79.5451
2.6812 | -79.5394
2.6815 | -79.5344
2.6818 | -79.5288
2.6821 | -79.5224
2.6825 | -79.5165
2.6828 | -79.51
2.6831 | -79.5051
2.6835 | -79.4982
2.6838 | -79.4924
2.6841 | -79.4873
2.6844 | -79.4784
2.6848 | -79.4717
2.6851 | -79.4662
2.6854 | -79.4599
2.6858 | -79.4535
2.6861 | -79.4466
2.6864 | -79.4395
2.6868 | -79.4332
2.6871 | -79.4272
2.6874 | -79.4214
2.6877 | -79.4147
2.6881 | -79.4066
2.6884 | -79.4004
2.6887 | -79.3954
2.6891 | -79.3897
2.6894 | -79.3837
2.6897 | -79.3757
2.69 | -79.3676
2.6904 | -79.36
2.6907 | -79.3511
2.691 | -79.3441
2.6914 | -79.3382
2.6917 | -79.3319
2.692 | -79.3223
2.6923 | -79.3159
2.6927 | -79.3093
2.693 | -79.3032
2.6933 | -79.2971
2.6937 | -79.2899
2.694 | -79.2833
2.6943 | -79.2771
2.6947 | -79.2706
2.695 | -79.2644
2.6953 | -79.2581
2.6956 | -79.2498
2.696 | -79.2438
2.6963 | -79.2376
2.6966 | -79.2315
2.697 | -79.2267
2.6973 | -79.2188
2.6976 | -79.2139
2.6979 | -79.2065
2.6983 | -79.1967
2.6986 | -79.1884
2.6989 | -79.183
2.6993 | -79.178
2.6996 | -79.1722
2.6999 | -79.1663
2.7002 | -79.16
2.7006 | -79.153
2.7009 | -79.1434
2.7012 | -79.1367
2.7016 | -79.1314
2.7019 | -79.1262
2.7022 | -79.1201
2.7026 | -79.114
2.7029 | -79.107
2.7032 | -79.0996
2.7035 | -79.0926
2.7039 | -79.0868
2.7042 | -79.0813
2.7045 | -79.0744
2.7049 | -79.0684
2.7052 | -79.0628
2.7055 | -79.0573
2.7058 | -79.0509
2.7062 | -79.0442
2.7065 | -79.0363
2.7068 | -79.0294
2.7072 | -79.0236
2.7075 | -79.019
2.7078 | -79.0135
2.7081 | -79.008
2.7085 | -79.0024
2.7088 | -78.9967
2.7091 | -78.9919
2.7095 | -78.9862
2.7098 | -78.982
2.7101 | -78.9776
2.7105 | -78.9724
2.7108 | -78.9666
2.7111 | -78.9623
2.7114 | -78.957
2.7118 | -78.9515
2.7121 | -78.9462
2.7124 | -78.941
2.7128 | -78.9349
2.7131 | -78.9299
2.7134 | -78.9253
2.7137 | -78.9215
2.7141 | -78.9165
2.7144 | -78.9134
2.7147 | -78.9099
2.7151 | -78.9043
2.7154 | -78.8997
2.7157 | -78.895
2.716 | -78.892
2.7164 | -78.8876
2.7167 | -78.884
2.717 | -78.879
2.7174 | -78.8744
2.7177 | -78.8702
2.718 | -78.8651
2.7184 | -78.8574
2.7187 | -78.8487
2.719 | -78.8436
2.7193 | -78.8397
2.7197 | -78.8354
2.72 | -78.8312
2.7203 | -78.8273
2.7207 | -78.8233
2.721 | -78.8186
2.7213 | -78.8146
2.7216 | -78.8114
2.722 | -78.8094
2.7223 | -78.8081
2.7226 | -78.806
2.723 | -78.8023
2.7233 | -78.7966
2.7236 | -78.793
2.7239 | -78.7899
2.7243 | -78.7854
2.7246 | -78.7812
2.7249 | -78.7774
2.7253 | -78.7733
2.7256 | -78.7697
2.7259 | -78.766
2.7262 | -78.7624
2.7266 | -78.7591
2.7269 | -78.7562
2.7272 | -78.7542
2.7276 | -78.7501
2.7279 | -78.747
2.7282 | -78.7432
2.7286 | -78.741
2.7289 | -78.7367
2.7292 | -78.7312
2.7295 | -78.7258
2.7299 | -78.7217
2.7302 | -78.7186
2.7305 | -78.7147
2.7309 | -78.7122
2.7312 | -78.7086
2.7315 | -78.7066
2.7318 | -78.7039
2.7322 | -78.7023
2.7325 | -78.6995
2.7328 | -78.6961
2.7332 | -78.6937
2.7335 | -78.6907
2.7338 | -78.6877
2.7341 | -78.6851
2.7345 | -78.6844
2.7348 | -78.6854
2.7351 | -78.6859
2.7355 | -78.6837
2.7358 | -78.6816
2.7361 | -78.6795
2.7365 | -78.6769
2.7368 | -78.6741
2.7371 | -78.6714
2.7374 | -78.6681
2.7378 | -78.6636
2.7381 | -78.6613
2.7384 | -78.6571
2.7388 | -78.6542
2.7391 | -78.6525
2.7394 | -78.6483
2.7397 | -78.6455
2.7401 | -78.644
2.7404 | -78.6424
2.7407 | -78.6414
2.7411 | -78.6399
2.7414 | -78.6378
2.7417 | -78.6349
2.742 | -78.6338
2.7424 | -78.6319
2.7427 | -78.6311
2.743 | -78.6296
2.7434 | -78.6285
2.7437 | -78.6274
2.744 | -78.6267
2.7444 | -78.6287
2.7447 | -78.6284
2.745 | -78.6271
2.7453 | -78.6259
2.7457 | -78.6253
2.746 | -78.6251
2.7463 | -78.6235
2.7467 | -78.6222
2.747 | -78.6214
2.7473 | -78.6202
2.7476 | -78.617
2.748 | -78.6142
2.7483 | -78.6107
2.7486 | -78.6061
2.749 | -78.6047
2.7493 | -78.6046
2.7496 | -78.6043
2.7499 | -78.6024
2.7503 | -78.5993
2.7506 | -78.5979
2.7509 | -78.5961
2.7513 | -78.5936
2.7516 | -78.593
2.7519 | -78.5929
2.7523 | -78.5921
2.7526 | -78.5893
2.7529 | -78.5893
2.7532 | -78.5882
2.7536 | -78.5868
2.7539 | -78.5858
2.7542 | -78.5862
2.7546 | -78.5868
2.7549 | -78.5865
2.7552 | -78.5859
2.7555 | -78.5838
2.7559 | -78.5824
2.7562 | -78.5819
2.7565 | -78.582
2.7569 | -78.5816
2.7572 | -78.5814
2.7575 | -78.5812
2.7578 | -78.5808
2.7582 | -78.5811
2.7585 | -78.582
2.7588 | -78.5829
2.7592 | -78.5845
2.7595 | -78.5847
2.7598 | -78.5827
2.7602 | -78.5814
2.7605 | -78.5813
2.7608 | -78.5819
2.7611 | -78.5807
2.7615 | -78.5775
2.7618 | -78.5778
2.7621 | -78.5786
2.7625 | -78.5749
2.7628 | -78.5734
2.7631 | -78.5733
2.7634 | -78.5753
2.7638 | -78.5753
2.7641 | -78.5751
2.7644 | -78.5744
2.7648 | -78.5735
2.7651 | -78.5738
2.7654 | -78.5751
2.7657 | -78.5748
2.7661 | -78.5769
2.7664 | -78.5772
2.7667 | -78.5775
2.7671 | -78.5779
2.7674 | -78.577
2.7677 | -78.575
2.7681 | -78.5704
2.7684 | -78.5654
2.7687 | -78.5652
2.769 | -78.5656
2.7694 | -78.5658
2.7697 | -78.5661
2.77 | -78.5663
2.7704 | -78.5665
2.7707 | -78.5672
2.771 | -78.5678
2.7713 | -78.5684
2.7717 | -78.569
2.772 | -78.5687
2.7723 | -78.5677
2.7727 | -78.5669
2.773 | -78.5677
2.7733 | -78.5687
2.7736 | -78.5678
2.774 | -78.5696
2.7743 | -78.5685
2.7746 | -78.5698
2.775 | -78.57
2.7753 | -78.5681
2.7756 | -78.5674
2.7759 | -78.5668
2.7763 | -78.5661
2.7766 | -78.5666
2.7769 | -78.5664
2.7773 | -78.5647
2.7776 | -78.5638
2.7779 | -78.5616
2.7783 | -78.5612
2.7786 | -78.5613
2.7789 | -78.5616
2.7792 | -78.5614
2.7796 | -78.5616
2.7799 | -78.5606
2.7802 | -78.5591
2.7806 | -78.5581
2.7809 | -78.5584
2.7812 | -78.5593
2.7815 | -78.559
2.7819 | -78.5604
2.7822 | -78.5591
2.7825 | -78.5602
2.7829 | -78.5605
2.7832 | -78.5604
2.7835 | -78.5617
2.7838 | -78.5611
2.7842 | -78.5606
2.7845 | -78.5603
2.7848 | -78.5582
2.7852 | -78.5548
2.7855 | -78.5532
2.7858 | -78.5528
2.7862 | -78.553
2.7865 | -78.5531
2.7868 | -78.5536
2.7871 | -78.5539
2.7875 | -78.5529
2.7878 | -78.5539
2.7881 | -78.5541
2.7885 | -78.5539
2.7888 | -78.5537
2.7891 | -78.5542
2.7894 | -78.5534
2.7898 | -78.5522
2.7901 | -78.5522
2.7904 | -78.5532
2.7908 | -78.554
2.7911 | -78.5548
2.7914 | -78.5552
2.7917 | -78.5527
2.7921 | -78.5542
2.7924 | -78.5542
2.7927 | -78.5541
2.7931 | -78.5541
2.7934 | -78.5538
2.7937 | -78.554
2.7941 | -78.5541
2.7944 | -78.5524
2.7947 | -78.5497
2.795 | -78.5479
2.7954 | -78.547
2.7957 | -78.5467
2.796 | -78.546
2.7964 | -78.5454
2.7967 | -78.5466
2.797 | -78.5489
2.7973 | -78.5494
2.7977 | -78.5493
2.798 | -78.5502
2.7983 | -78.5492
2.7987 | -78.5516
2.799 | -78.5541
2.7993 | -78.5557
2.7996 | -78.5533
2.8 | -78.554
2.8003 | -78.5536
2.8006 | -78.5533
2.801 | -78.5534
2.8013 | -78.5506
2.8016 | -78.5494
2.802 | -78.5487
2.8023 | -78.5471
2.8026 | -78.5442
2.8029 | -78.5429
2.8033 | -78.5425
2.8036 | -78.5421
2.8039 | -78.5416
2.8043 | -78.5409
2.8046 | -78.5396
2.8049 | -78.5391
2.8052 | -78.5373
2.8056 | -78.5343
2.8059 | -78.5331
2.8062 | -78.5345
2.8066 | -78.5341
2.8069 | -78.5319
2.8072 | -78.5299
2.8075 | -78.5288
2.8079 | -78.5282
2.8082 | -78.5278
2.8085 | -78.5267
2.8089 | -78.5248
2.8092 | -78.5228
2.8095 | -78.5205
2.8098 | -78.5193
2.8102 | -78.5185
2.8105 | -78.5165
2.8108 | -78.5128
2.8112 | -78.5103
2.8115 | -78.5102
2.8118 | -78.5093
2.8122 | -78.5077
2.8125 | -78.5066
2.8128 | -78.5061
2.8131 | -78.5051
2.8135 | -78.5038
2.8138 | -78.5034
2.8141 | -78.5038
2.8145 | -78.5029
2.8148 | -78.502
2.8151 | -78.5009
2.8154 | -78.4999
2.8158 | -78.4995
2.8161 | -78.4992
2.8164 | -78.5011
2.8168 | -78.4984
2.8171 | -78.4939
2.8174 | -78.4936
2.8178 | -78.4928
2.8181 | -78.4896
2.8184 | -78.4874
2.8187 | -78.486
2.8191 | -78.4833
2.8194 | -78.48
2.8197 | -78.4773
2.8201 | -78.4754
2.8204 | -78.4741
2.8207 | -78.4728
2.821 | -78.4721
2.8214 | -78.4713
2.8217 | -78.47
2.822 | -78.4675
2.8224 | -78.465
2.8227 | -78.4631
2.823 | -78.4609
2.8233 | -78.4583
2.8237 | -78.4579
2.824 | -78.4564
2.8243 | -78.4548
2.8247 | -78.4538
2.825 | -78.4533
2.8253 | -78.4512
2.8256 | -78.4495
2.826 | -78.4471
2.8263 | -78.4458
2.8266 | -78.4442
2.827 | -78.4421
2.8273 | -78.4403
2.8276 | -78.4392
2.828 | -78.44
2.8283 | -78.4396
2.8286 | -78.4381
2.8289 | -78.4359
2.8293 | -78.4343
2.8296 | -78.4327
2.8299 | -78.4302
2.8303 | -78.4282
2.8306 | -78.4258
2.8309 | -78.424
2.8312 | -78.4243
2.8316 | -78.4237
2.8319 | -78.4207
2.8322 | -78.4183
2.8326 | -78.4172
2.8329 | -78.4159
2.8332 | -78.4143
2.8335 | -78.4134
2.8339 | -78.4112
2.8342 | -78.4091
2.8345 | -78.4076
2.8349 | -78.4078
2.8352 | -78.4071
2.8355 | -78.407
2.8359 | -78.4054
2.8362 | -78.4029
2.8365 | -78.4016
2.8368 | -78.3998
2.8372 | -78.3993
2.8375 | -78.3986
2.8378 | -78.3972
2.8382 | -78.3953
2.8385 | -78.3941
2.8388 | -78.3938
2.8391 | -78.3917
2.8395 | -78.3909
2.8398 | -78.3906
2.8401 | -78.3882
2.8405 | -78.3881
2.8408 | -78.3874
2.8411 | -78.3875
2.8414 | -78.3863
2.8418 | -78.3842
2.8421 | -78.3841
2.8424 | -78.3839
2.8428 | -78.3827
2.8431 | -78.3803
2.8434 | -78.3786
2.8438 | -78.3772
2.8441 | -78.3756
2.8444 | -78.3749
2.8447 | -78.374
2.8451 | -78.3716
2.8454 | -78.368
2.8457 | -78.3662
2.8461 | -78.3645
2.8464 | -78.3635
2.8467 | -78.3616
2.847 | -78.3579
2.8474 | -78.3556
2.8477 | -78.3565
2.848 | -78.3554
2.8484 | -78.3524
2.8487 | -78.3511
2.849 | -78.3478
2.8493 | -78.346
2.8497 | -78.344
2.85 | -78.3423
2.8503 | -78.3421
2.8507 | -78.341
2.851 | -78.3417
2.8513 | -78.3415
2.8517 | -78.3441
2.852 | -78.3452
2.8523 | -78.3439
2.8526 | -78.3429
2.853 | -78.343
2.8533 | -78.3432
2.8536 | -78.3411
2.854 | -78.3408
2.8543 | -78.3392
2.8546 | -78.3383
2.8549 | -78.3343
2.8553 | -78.3335
2.8556 | -78.3345
2.8559 | -78.3362
2.8563 | -78.3351
2.8566 | -78.3348
2.8569 | -78.3339
2.8572 | -78.3325
2.8576 | -78.3303
2.8579 | -78.3284
2.8582 | -78.329
2.8586 | -78.329
2.8589 | -78.3298
2.8592 | -78.3289
2.8596 | -78.3283
2.8599 | -78.3266
2.8602 | -78.3257
2.8605 | -78.3246
2.8609 | -78.3228
2.8612 | -78.3219
2.8615 | -78.3211
2.8619 | -78.3205
2.8622 | -78.3205
2.8625 | -78.3211
2.8628 | -78.3189
2.8632 | -78.3162
2.8635 | -78.3148
2.8638 | -78.3141
2.8642 | -78.3133
2.8645 | -78.3136
2.8648 | -78.3129
2.8651 | -78.3132
2.8655 | -78.3131
2.8658 | -78.313
2.8661 | -78.3127
2.8665 | -78.3117
2.8668 | -78.3109
2.8671 | -78.3104
2.8674 | -78.31
2.8678 | -78.3099
2.8681 | -78.3106
2.8684 | -78.311
2.8688 | -78.3108
2.8691 | -78.3106
2.8694 | -78.3096
2.8698 | -78.3096
2.8701 | -78.3093
2.8704 | -78.309
2.8707 | -78.3089
2.8711 | -78.3084
2.8714 | -78.3078
2.8717 | -78.3077
2.8721 | -78.3073
2.8724 | -78.3082
2.8727 | -78.3077
2.873 | -78.3066
2.8734 | -78.3068
2.8737 | -78.3071
2.874 | -78.3071
2.8744 | -78.307
2.8747 | -78.3071
2.875 | -78.3071
2.8753 | -78.3068
2.8757 | -78.3055
2.876 | -78.3052
2.8763 | -78.3052
2.8767 | -78.3065
2.877 | -78.3075
2.8773 | -78.3102
2.8777 | -78.3102
2.878 | -78.311
2.8783 | -78.3114
2.8786 | -78.3125
2.879 | -78.3131
2.8793 | -78.3138
2.8796 | -78.3139
2.88 | -78.314
2.8803 | -78.3146
2.8806 | -78.3149
2.8809 | -78.3148
2.8813 | -78.3144
2.8816 | -78.3147
2.8819 | -78.3166
2.8823 | -78.3174
2.8826 | -78.3183
2.8829 | -78.3186
2.8832 | -78.3184
2.8836 | -78.3207
2.8839 | -78.3228
2.8842 | -78.3243
2.8846 | -78.3253
2.8849 | -78.3258
2.8852 | -78.3268
2.8856 | -78.3289
2.8859 | -78.3322
2.8862 | -78.3337
2.8865 | -78.3353
2.8869 | -78.3375
2.8872 | -78.3385
2.8875 | -78.3392
2.8879 | -78.3403
2.8882 | -78.3414
2.8885 | -78.3406
2.8888 | -78.3418
2.8892 | -78.3427
2.8895 | -78.343
2.8898 | -78.3446
2.8902 | -78.3453
2.8905 | -78.344
2.8908 | -78.3449
2.8911 | -78.346
2.8915 | -78.3477
2.8918 | -78.3502
2.8921 | -78.3509
2.8925 | -78.354
2.8928 | -78.3583
2.8931 | -78.3609
2.8935 | -78.3618
2.8938 | -78.363
2.8941 | -78.3634
2.8944 | -78.3627
2.8948 | -78.3625
2.8951 | -78.3635
2.8954 | -78.3669
2.8958 | -78.3678
2.8961 | -78.3658
2.8964 | -78.3634
2.8967 | -78.3625
2.8971 | -78.3634
2.8974 | -78.3661
2.8977 | -78.368
2.8981 | -78.3702
2.8984 | -78.3725
2.8987 | -78.3712
2.899 | -78.3694
2.8994 | -78.3713
2.8997 | -78.3718
2.9 | -78.3702
2.9004 | -78.369
2.9007 | -78.3701
2.901 | -78.3713
2.9014 | -78.3729
2.9017 | -78.3737
2.902 | -78.3749
2.9023 | -78.3769
2.9027 | -78.3779
2.903 | -78.3789
2.9033 | -78.38
2.9037 | -78.381
2.904 | -78.3811
2.9043 | -78.3818
2.9046 | -78.3829
2.905 | -78.3833
2.9053 | -78.3845
2.9056 | -78.3869
2.906 | -78.3882
2.9063 | -78.3894
2.9066 | -78.389
2.9069 | -78.389
2.9073 | -78.3904
2.9076 | -78.3923
2.9079 | -78.394
2.9083 | -78.3956
2.9086 | -78.3967
2.9089 | -78.3979
2.9093 | -78.3987
2.9096 | -78.3995
2.9099 | -78.3982
2.9102 | -78.3975
2.9106 | -78.3993
2.9109 | -78.4014
2.9112 | -78.4028
2.9116 | -78.4031
2.9119 | -78.4027
2.9122 | -78.4037
2.9125 | -78.4042
2.9129 | -78.4042
2.9132 | -78.405
2.9135 | -78.4067
2.9139 | -78.408
2.9142 | -78.4126
2.9145 | -78.4156
2.9148 | -78.4134
2.9152 | -78.4128
2.9155 | -78.414
2.9158 | -78.4153
2.9162 | -78.4152
2.9165 | -78.4141
2.9168 | -78.4143
2.9171 | -78.4156
2.9175 | -78.4168
2.9178 | -78.4179
2.9181 | -78.4173
2.9185 | -78.4156
2.9188 | -78.4161
2.9191 | -78.4166
2.9195 | -78.4184
2.9198 | -78.4198
2.9201 | -78.4208
2.9204 | -78.423
2.9208 | -78.4245
2.9211 | -78.4253
2.9214 | -78.428
2.9218 | -78.43
2.9221 | -78.431
2.9224 | -78.4316
2.9227 | -78.4335
2.9231 | -78.4334
2.9234 | -78.4341
2.9237 | -78.4354
2.9241 | -78.4378
2.9244 | -78.4384
2.9247 | -78.4381
2.925 | -78.4394
2.9254 | -78.4384
2.9257 | -78.4397
2.926 | -78.4421
2.9264 | -78.4435
2.9267 | -78.4448
2.927 | -78.4454
2.9274 | -78.4467
2.9277 | -78.4475
2.928 | -78.4485
2.9283 | -78.4493
2.9287 | -78.45
2.929 | -78.4501
2.9293 | -78.4506
2.9297 | -78.4523
2.93 | -78.453
2.9303 | -78.4525
2.9306 | -78.4523
2.931 | -78.4525
2.9313 | -78.4495
2.9316 | -78.4478
2.932 | -78.4484
2.9323 | -78.4487
2.9326 | -78.4487
2.9329 | -78.4477
2.9333 | -78.445
2.9336 | -78.4451
2.9339 | -78.4464
2.9343 | -78.4464
2.9346 | -78.4456
2.9349 | -78.4463
2.9353 | -78.4486
2.9356 | -78.4463
2.9359 | -78.4478
2.9362 | -78.4494
2.9366 | -78.4496
2.9369 | -78.4516
2.9372 | -78.4539
2.9376 | -78.4549
2.9379 | -78.4548
2.9382 | -78.4545
2.9385 | -78.454
2.9389 | -78.4537
2.9392 | -78.4549
2.9395 | -78.4551
2.9399 | -78.4548
2.9402 | -78.4547
2.9405 | -78.4548
2.9408 | -78.4547
2.9412 | -78.4527
2.9415 | -78.4481
2.9418 | -78.4417
2.9422 | -78.438
2.9425 | -78.4363
2.9428 | -78.4348
2.9432 | -78.4335
2.9435 | -78.4331
2.9438 | -78.4321
2.9441 | -78.4311
2.9445 | -78.431
2.9448 | -78.4284
2.9451 | -78.4239
2.9455 | -78.4212
2.9458 | -78.4229
2.9461 | -78.4244
2.9464 | -78.426
2.9468 | -78.4284
2.9471 | -78.4284
2.9474 | -78.4267
2.9478 | -78.425
2.9481 | -78.4252
2.9484 | -78.4274
2.9487 | -78.4271
2.9491 | -78.4268
2.9494 | -78.4272
2.9497 | -78.425
2.9501 | -78.4236
2.9504 | -78.4229
2.9507 | -78.4237
2.9511 | -78.4235
2.9514 | -78.4232
2.9517 | -78.423
2.952 | -78.4219
2.9524 | -78.4219
2.9527 | -78.421
2.953 | -78.4195
2.9534 | -78.4163
2.9537 | -78.4134
2.954 | -78.4125
2.9543 | -78.4128
2.9547 | -78.4118
2.955 | -78.4082
2.9553 | -78.4036
2.9557 | -78.4024
2.956 | -78.4013
2.9563 | -78.398
2.9566 | -78.3935
2.957 | -78.3897
2.9573 | -78.3864
2.9576 | -78.3855
2.958 | -78.3836
2.9583 | -78.3818
2.9586 | -78.382
2.959 | -78.3803
2.9593 | -78.3781
2.9596 | -78.3764
2.9599 | -78.3764
2.9603 | -78.3777
2.9606 | -78.3755
2.9609 | -78.3719
2.9613 | -78.3731
2.9616 | -78.3724
2.9619 | -78.37
2.9622 | -78.3663
2.9626 | -78.363
2.9629 | -78.3599
2.9632 | -78.3574
2.9636 | -78.3554
2.9639 | -78.3535
2.9642 | -78.3516
2.9645 | -78.3497
2.9649 | -78.3477
2.9652 | -78.346
2.9655 | -78.342
2.9659 | -78.3394
2.9662 | -78.3367
2.9665 | -78.3346
2.9668 | -78.3338
2.9672 | -78.3338
2.9675 | -78.3369
2.9678 | -78.3384
2.9682 | -78.3331
2.9685 | -78.3298
2.9688 | -78.3274
2.9692 | -78.3242
2.9695 | -78.3219
2.9698 | -78.3195
2.9701 | -78.317
2.9705 | -78.3146
2.9708 | -78.311
2.9711 | -78.3063
2.9715 | -78.3043
2.9718 | -78.3037
2.9721 | -78.3001
2.9724 | -78.2954
2.9728 | -78.2903
2.9731 | -78.2882
2.9734 | -78.2856
2.9738 | -78.2832
2.9741 | -78.2805
2.9744 | -78.2786
2.9747 | -78.2727
2.9751 | -78.2689
2.9754 | -78.2682
2.9757 | -78.2665
2.9761 | -78.2646
2.9764 | -78.2623
2.9767 | -78.2585
2.9771 | -78.2559
2.9774 | -78.2522
2.9777 | -78.2502
2.978 | -78.249
2.9784 | -78.2456
2.9787 | -78.2416
2.979 | -78.238
2.9794 | -78.2348
2.9797 | -78.23
2.98 | -78.227
2.9803 | -78.2246
2.9807 | -78.2228
2.981 | -78.2192
2.9813 | -78.2161
2.9817 | -78.2118
2.982 | -78.2091
2.9823 | -78.2066
2.9826 | -78.2046
2.983 | -78.2013
2.9833 | -78.1975
2.9836 | -78.1943
2.984 | -78.1904
2.9843 | -78.1873
2.9846 | -78.1844
2.985 | -78.1819
2.9853 | -78.1796
2.9856 | -78.1783
2.9859 | -78.178
2.9863 | -78.1765
2.9866 | -78.1739
2.9869 | -78.1706
2.9873 | -78.1677
2.9876 | -78.1647
2.9879 | -78.1616
2.9882 | -78.1598
2.9886 | -78.1569
2.9889 | -78.1535
2.9892 | -78.1505
2.9896 | -78.1476
2.9899 | -78.1445
2.9902 | -78.1422
2.9905 | -78.1398
2.9909 | -78.1368
2.9912 | -78.1343
2.9915 | -78.131
2.9919 | -78.1275
2.9922 | -78.1236
2.9925 | -78.1213
2.9929 | -78.1202
2.9932 | -78.1172
2.9935 | -78.1144
2.9938 | -78.1117
2.9942 | -78.109
2.9945 | -78.1048
2.9948 | -78.1017
2.9952 | -78.099
2.9955 | -78.0971
2.9958 | -78.0961
2.9961 | -78.0956
2.9965 | -78.0942
2.9968 | -78.0914
2.9971 | -78.0876
2.9975 | -78.0845
2.9978 | -78.081
2.9981 | -78.0776
2.9984 | -78.0749
2.9988 | -78.0727
2.9991 | -78.071
2.9994 | -78.0687
2.9998 | -78.0659
3.0001 | -78.0634
3.0004 | -78.0614
3.0008 | -78.0612
3.0011 | -78.0603
3.0014 | -78.0595
3.0017 | -78.0577
3.0021 | -78.0555
3.0024 | -78.0532
3.0027 | -78.0514
3.0031 | -78.0505
3.0034 | -78.0493
3.0037 | -78.0473
3.004 | -78.0452
3.0044 | -78.044
3.0047 | -78.0439
3.005 | -78.0419
3.0054 | -78.0398
3.0057 | -78.038
3.006 | -78.0365
3.0063 | -78.0346
3.0067 | -78.0328
3.007 | -78.0309
3.0073 | -78.0297
3.0077 | -78.0302
3.008 | -78.028
3.0083 | -78.0247
3.0086 | -78.0228
3.009 | -78.0226
3.0093 | -78.0212
3.0096 | -78.0193
3.01 | -78.0177
3.0103 | -78.0181
3.0106 | -78.0157
3.011 | -78.0137
3.0113 | -78.013
3.0116 | -78.0123
3.0119 | -78.012
3.0123 | -78.0109
3.0126 | -78.0096
3.0129 | -78.0092
3.0133 | -78.0094
3.0136 | -78.0089
3.0139 | -78.0087
3.0142 | -78.0076
3.0146 | -78.006
3.0149 | -78.0055
3.0152 | -78.0047
3.0156 | -78.0021
3.0159 | -78.0012
3.0162 | -78.0004
3.0165 | -77.9993
3.0169 | -77.9981
3.0172 | -77.9959
3.0175 | -77.9951
3.0179 | -77.9945
3.0182 | -77.9939
3.0185 | -77.9932
3.0189 | -77.9909
3.0192 | -77.9869
3.0195 | -77.9875
3.0198 | -77.9878
3.0202 | -77.9903
3.0205 | -77.9891
3.0208 | -77.9881
3.0212 | -77.9879
3.0215 | -77.9869
3.0218 | -77.9858
3.0221 | -77.9861
3.0225 | -77.9874
3.0228 | -77.9872
3.0231 | -77.9879
3.0235 | -77.986
3.0238 | -77.9851
3.0241 | -77.985
3.0244 | -77.9846
3.0248 | -77.9843
3.0251 | -77.9836
3.0254 | -77.9845
3.0258 | -77.9852
3.0261 | -77.9858
3.0264 | -77.9872
3.0268 | -77.988
3.0271 | -77.9882
3.0274 | -77.9883
3.0277 | -77.9875
3.0281 | -77.9863
3.0284 | -77.9849
3.0287 | -77.9837
3.0291 | -77.9861
3.0294 | -77.9858
3.0297 | -77.9858
3.03 | -77.986
3.0304 | -77.9864
3.0307 | -77.9871
3.031 | -77.9876
3.0314 | -77.9866
3.0317 | -77.9865
3.032 | -77.9874
3.0323 | -77.9879
3.0327 | -77.9888
3.033 | -77.9894
3.0333 | -77.9898
3.0337 | -77.9918
3.034 | -77.9931
3.0343 | -77.9947
3.0347 | -77.9965
3.035 | -77.9981
3.0353 | -77.9987
3.0356 | -77.9978
3.036 | -77.9985
3.0363 | -78.0003
3.0366 | -78.0011
3.037 | -78.0019
3.0373 | -78.0031
3.0376 | -78.004
3.0379 | -78.006
3.0383 | -78.0082
3.0386 | -78.011
3.0389 | -78.0129
3.0393 | -78.0141
3.0396 | -78.0148
3.0399 | -78.0151
3.0402 | -78.0161
3.0406 | -78.0184
3.0409 | -78.0194
3.0412 | -78.0206
3.0416 | -78.0213
3.0419 | -78.0224
3.0422 | -78.024
3.0426 | -78.0265
3.0429 | -78.0285
3.0432 | -78.031
3.0435 | -78.0339
3.0439 | -78.035
3.0442 | -78.037
3.0445 | -78.0392
3.0449 | -78.042
3.0452 | -78.0434
3.0455 | -78.0456
3.0458 | -78.0471
3.0462 | -78.0485
3.0465 | -78.0519
3.0468 | -78.0542
3.0472 | -78.0567
3.0475 | -78.0594
3.0478 | -78.0631
3.0481 | -78.0686
3.0485 | -78.0715
3.0488 | -78.0744
3.0491 | -78.0761
3.0495 | -78.0762
3.0498 | -78.0754
3.0501 | -78.0762
3.0505 | -78.0803
3.0508 | -78.0842
3.0511 | -78.087
3.0514 | -78.0885
3.0518 | -78.0921
3.0521 | -78.0972
3.0524 | -78.101
3.0528 | -78.1045
3.0531 | -78.1085
3.0534 | -78.1127
3.0537 | -78.1153
3.0541 | -78.1172
3.0544 | -78.1209
3.0547 | -78.1249
3.0551 | -78.1297
3.0554 | -78.1331
3.0557 | -78.1362
3.056 | -78.1393
3.0564 | -78.1421
3.0567 | -78.1434
3.057 | -78.1452
3.0574 | -78.1482
3.0577 | -78.1531
3.058 | -78.1565
3.0583 | -78.1596
3.0587 | -78.1625
3.059 | -78.1653
3.0593 | -78.1689
3.0597 | -78.1719
3.06 | -78.1756
3.0603 | -78.1791
3.0607 | -78.1851
3.061 | -78.1903
3.0613 | -78.1937
3.0616 | -78.1968
3.062 | -78.2002
3.0623 | -78.2035
3.0626 | -78.2071
3.063 | -78.2112
3.0633 | -78.2163
3.0636 | -78.2203
3.0639 | -78.2234
3.0643 | -78.2275
3.0646 | -78.2307
3.0649 | -78.2338
3.0653 | -78.2377
3.0656 | -78.2408
3.0659 | -78.2449
3.0662 | -78.249
3.0666 | -78.2514
3.0669 | -78.2561
3.0672 | -78.2617
3.0676 | -78.2655
3.0679 | -78.2693
3.0682 | -78.2723
3.0686 | -78.2758
3.0689 | -78.2793
3.0692 | -78.2848
3.0695 | -78.2901
3.0699 | -78.2942
3.0702 | -78.2965
3.0705 | -78.2992
3.0709 | -78.3034
3.0712 | -78.307
3.0715 | -78.3104
3.0718 | -78.3144
3.0722 | -78.3187
3.0725 | -78.3229
3.0728 | -78.3277
3.0732 | -78.3321
3.0735 | -78.3358
3.0738 | -78.3402
3.0741 | -78.3458
3.0745 | -78.3504
3.0748 | -78.3545
3.0751 | -78.359
3.0755 | -78.3632
3.0758 | -78.3675
3.0761 | -78.3723
3.0765 | -78.3762
3.0768 | -78.3802
3.0771 | -78.3847
3.0774 | -78.3893
3.0778 | -78.3941
3.0781 | -78.3986
3.0784 | -78.4034
3.0788 | -78.4073
3.0791 | -78.4099
3.0794 | -78.413
3.0797 | -78.4176
3.0801 | -78.4231
3.0804 | -78.4278
3.0807 | -78.4326
3.0811 | -78.4384
3.0814 | -78.443
3.0817 | -78.4474
3.082 | -78.4521
3.0824 | -78.4574
3.0827 | -78.4632
3.083 | -78.4684
3.0834 | -78.4736
3.0837 | -78.4784
3.084 | -78.4826
3.0844 | -78.4869
3.0847 | -78.4915
3.085 | -78.4966
3.0853 | -78.5025
3.0857 | -78.5077
3.086 | -78.5122
3.0863 | -78.5174
3.0867 | -78.521
3.087 | -78.5251
3.0873 | -78.5306
3.0876 | -78.5358
3.088 | -78.5405
3.0883 | -78.5459
3.0886 | -78.5516
3.089 | -78.5573
3.0893 | -78.563
3.0896 | -78.567
3.0899 | -78.5707
3.0903 | -78.5753
3.0906 | -78.5806
3.0909 | -78.5855
3.0913 | -78.5898
3.0916 | -78.5937
3.0919 | -78.5976
3.0923 | -78.6041
3.0926 | -78.6099
3.0929 | -78.6142
3.0932 | -78.619
3.0936 | -78.6228
3.0939 | -78.6267
3.0942 | -78.6313
3.0946 | -78.6362
3.0949 | -78.6407
3.0952 | -78.6445
3.0955 | -78.6486
3.0959 | -78.6524
3.0962 | -78.6581
3.0965 | -78.6625
3.0969 | -78.6689
3.0972 | -78.6737
3.0975 | -78.6784
3.0978 | -78.6841
3.0982 | -78.6895
3.0985 | -78.6944
3.0988 | -78.7007
3.0992 | -78.7045
3.0995 | -78.7086
3.0998 | -78.7118
3.1002 | -78.7157
3.1005 | -78.7197
3.1008 | -78.7232
3.1011 | -78.7263
3.1015 | -78.7301
3.1018 | -78.7343
3.1021 | -78.7386
3.1025 | -78.744
3.1028 | -78.7496
3.1031 | -78.7551
3.1034 | -78.7604
3.1038 | -78.7653
3.1041 | -78.7705
3.1044 | -78.7767
3.1048 | -78.784
3.1051 | -78.7871
3.1054 | -78.7949
3.1057 | -78.8009
3.1061 | -78.8055
3.1064 | -78.8092
3.1067 | -78.8134
3.1071 | -78.8175
3.1074 | -78.8216
3.1077 | -78.8277
3.108 | -78.8349
3.1084 | -78.8395
3.1087 | -78.8445
3.109 | -78.8512
3.1094 | -78.8572
3.1097 | -78.8609
3.11 | -78.8654
3.1104 | -78.8706
3.1107 | -78.8765
3.111 | -78.8815
3.1113 | -78.8843
3.1117 | -78.8877
3.112 | -78.8917
3.1123 | -78.8956
3.1127 | -78.8991
3.113 | -78.9039
3.1133 | -78.908
3.1136 | -78.9116
3.114 | -78.9157
3.1143 | -78.9193
3.1146 | -78.9232
3.115 | -78.9272
3.1153 | -78.9303
3.1156 | -78.9339
3.1159 | -78.9382
3.1163 | -78.9438
3.1166 | -78.9488
3.1169 | -78.9527
3.1173 | -78.9559
3.1176 | -78.9592
3.1179 | -78.9637
3.1183 | -78.9672
3.1186 | -78.97
3.1189 | -78.9742
3.1192 | -78.9792
3.1196 | -78.9848
3.1199 | -78.9896
3.1202 | -78.9926
3.1206 | -78.9967
3.1209 | -79.0017
3.1212 | -79.0075
3.1215 | -79.0115
3.1219 | -79.0137
3.1222 | -79.0174
3.1225 | -79.0227
3.1229 | -79.0263
3.1232 | -79.0296
3.1235 | -79.0337
3.1238 | -79.037
3.1242 | -79.041
3.1245 | -79.0464
3.1248 | -79.0542
3.1252 | -79.0594
3.1255 | -79.0627
3.1258 | -79.066
3.1262 | -79.0692
3.1265 | -79.0739
3.1268 | -79.0799
3.1271 | -79.0842
3.1275 | -79.0872
3.1278 | -79.0888
3.1281 | -79.0912
3.1285 | -79.0963
3.1288 | -79.1027
3.1291 | -79.1066
3.1294 | -79.1109
3.1298 | -79.1147
3.1301 | -79.1171
3.1304 | -79.1209
3.1308 | -79.1243
3.1311 | -79.1274
3.1314 | -79.1318
3.1317 | -79.1353
3.1321 | -79.1381
3.1324 | -79.1411
3.1327 | -79.1449
3.1331 | -79.1482
3.1334 | -79.1524
3.1337 | -79.1554
3.1341 | -79.1587
3.1344 | -79.1608
3.1347 | -79.1642
3.135 | -79.1678
3.1354 | -79.1725
3.1357 | -79.1763
3.136 | -79.1799
3.1364 | -79.1846
3.1367 | -79.1872
3.137 | -79.1902
3.1373 | -79.1956
3.1377 | -79.1981
3.138 | -79.2014
3.1383 | -79.2047
3.1387 | -79.2083
3.139 | -79.2116
3.1393 | -79.215
3.1396 | -79.2181
3.14 | -79.2214
3.1403 | -79.2249
3.1406 | -79.2279
3.141 | -79.2318
3.1413 | -79.2374
3.1416 | -79.2417
3.142 | -79.2458
3.1423 | -79.2532
3.1426 | -79.2584
3.1429 | -79.2612
3.1433 | -79.2658
3.1436 | -79.2694
3.1439 | -79.2733
3.1443 | -79.2783
3.1446 | -79.2853
3.1449 | -79.2894
3.1452 | -79.2923
3.1456 | -79.2948
3.1459 | -79.2985
3.1462 | -79.3037
3.1466 | -79.3082
3.1469 | -79.3112
3.1472 | -79.3169
3.1475 | -79.3211
3.1479 | -79.3237
3.1482 | -79.3273
3.1485 | -79.3305
3.1489 | -79.3342
3.1492 | -79.3375
3.1495 | -79.3422
3.1498 | -79.3461
3.1502 | -79.3496
3.1505 | -79.3538
3.1508 | -79.3563
3.1512 | -79.3594
3.1515 | -79.3633
3.1518 | -79.3669
3.1522 | -79.3704
3.1525 | -79.3738
3.1528 | -79.3777
3.1531 | -79.3818
3.1535 | -79.3906
3.1538 | -79.4008
3.1541 | -79.4078
3.1545 | -79.4119
3.1548 | -79.4165
3.1551 | -79.4211
3.1554 | -79.4235
3.1558 | -79.4265
3.1561 | -79.43
3.1564 | -79.434
3.1568 | -79.4371
3.1571 | -79.4431
3.1574 | -79.4508
3.1577 | -79.4556
3.1581 | -79.4587
3.1584 | -79.4613
3.1587 | -79.4639
3.1591 | -79.4649
3.1594 | -79.4665
3.1597 | -79.4707
3.1601 | -79.475
3.1604 | -79.4791
3.1607 | -79.4829
3.161 | -79.4865
3.1614 | -79.4887
3.1617 | -79.491
3.162 | -79.4952
3.1624 | -79.4999
3.1627 | -79.5037
3.163 | -79.505
3.1633 | -79.5065
3.1637 | -79.5095
3.164 | -79.513
3.1643 | -79.5154
3.1647 | -79.5197
3.165 | -79.5233
3.1653 | -79.525
3.1656 | -79.5285
3.166 | -79.5315
3.1663 | -79.5361
3.1666 | -79.5413
3.167 | -79.5447
3.1673 | -79.5479
3.1676 | -79.5509
3.168 | -79.5561
3.1683 | -79.5618
3.1686 | -79.5647
3.1689 | -79.5669
3.1693 | -79.5717
3.1696 | -79.5787
3.1699 | -79.5839
3.1703 | -79.5893
3.1706 | -79.5927
3.1709 | -79.5961
3.1712 | -79.5996
3.1716 | -79.6046
3.1719 | -79.6079
3.1722 | -79.6127
3.1726 | -79.6168
3.1729 | -79.6199
3.1732 | -79.6227
3.1735 | -79.6259
3.1739 | -79.6301
3.1742 | -79.6353
3.1745 | -79.6389
3.1749 | -79.641
3.1752 | -79.6442
3.1755 | -79.6501
3.1759 | -79.6551
3.1762 | -79.6582
3.1765 | -79.6623
3.1768 | -79.6655
3.1772 | -79.6685
3.1775 | -79.6729
3.1778 | -79.6773
3.1782 | -79.6806
3.1785 | -79.684
3.1788 | -79.6879
3.1791 | -79.6933
3.1795 | -79.6975
3.1798 | -79.7018
3.1801 | -79.7057
3.1805 | -79.7079
3.1808 | -79.7109
3.1811 | -79.7158
3.1814 | -79.7196
3.1818 | -79.7234
3.1821 | -79.7307
3.1824 | -79.7351
3.1828 | -79.738
3.1831 | -79.7421
3.1834 | -79.7456
3.1838 | -79.7495
3.1841 | -79.753
3.1844 | -79.7563
3.1847 | -79.7614
3.1851 | -79.7682
3.1854 | -79.7719
3.1857 | -79.7752
3.1861 | -79.7776
3.1864 | -79.7818
3.1867 | -79.7902
3.187 | -79.7973
3.1874 | -79.8011
3.1877 | -79.8054
3.188 | -79.8094
3.1884 | -79.8115
3.1887 | -79.8139
3.189 | -79.8218
3.1893 | -79.8271
3.1897 | -79.8303
3.19 | -79.8333
3.1903 | -79.8366
3.1907 | -79.8403
3.191 | -79.8419
3.1913 | -79.8444
3.1917 | -79.8466
3.192 | -79.8528
3.1923 | -79.8574
3.1926 | -79.8613
3.193 | -79.8665
3.1933 | -79.8712
3.1936 | -79.8756
3.194 | -79.8796
3.1943 | -79.8855
3.1946 | -79.8895
3.1949 | -79.8936
3.1953 | -79.8966
3.1956 | -79.9002
3.1959 | -79.9032
3.1963 | -79.9076
3.1966 | -79.9128
3.1969 | -79.9166
3.1972 | -79.9198
3.1976 | -79.922
3.1979 | -79.9268
3.1982 | -79.9327
3.1986 | -79.9372
3.1989 | -79.941
3.1992 | -79.9446
3.1995 | -79.9485
3.1999 | -79.9522
3.2002 | -79.9563
3.2005 | -79.9589
3.2009 | -79.9608
3.2012 | -79.9631
3.2015 | -79.9659
3.2019 | -79.9692
3.2022 | -79.9731
3.2025 | -79.9765
3.2028 | -79.9797
3.2032 | -79.983
3.2035 | -79.9849
3.2038 | -79.9867
3.2042 | -79.9902
3.2045 | -79.9935
3.2048 | -79.9969
3.2051 | -80.0002
3.2055 | -80.0037
3.2058 | -80.0074
3.2061 | -80.0104
3.2065 | -80.0147
3.2068 | -80.0184
3.2071 | -80.0216
3.2074 | -80.0256
3.2078 | -80.0306
3.2081 | -80.0341
3.2084 | -80.0383
3.2088 | -80.0421
3.2091 | -80.0455
3.2094 | -80.0495
3.2098 | -80.053
3.2101 | -80.0558
3.2104 | -80.0608
3.2107 | -80.0648
3.2111 | -80.0689
3.2114 | -80.0725
3.2117 | -80.0761
3.2121 | -80.0803
3.2124 | -80.084
3.2127 | -80.0882
3.213 | -80.0936
3.2134 | -80.0982
3.2137 | -80.1032
3.214 | -80.1073
3.2144 | -80.1107
3.2147 | -80.1136
3.215 | -80.1172
3.2153 | -80.1202
3.2157 | -80.1242
3.216 | -80.127
3.2163 | -80.1288
3.2167 | -80.1319
3.217 | -80.1327
3.2173 | -80.1331
3.2177 | -80.1368
3.218 | -80.1403
3.2183 | -80.1437
3.2186 | -80.1477
3.219 | -80.152
3.2193 | -80.1535
3.2196 | -80.1542
3.22 | -80.1574
3.2203 | -80.1597
3.2206 | -80.1609
3.2209 | -80.1628
3.2213 | -80.1643
3.2216 | -80.1661
3.2219 | -80.1666
3.2223 | -80.1673
3.2226 | -80.1698
3.2229 | -80.1731
3.2232 | -80.1762
3.2236 | -80.1792
3.2239 | -80.1817
3.2242 | -80.1837
3.2246 | -80.1855
3.2249 | -80.1897
3.2252 | -80.1932
3.2256 | -80.1963
3.2259 | -80.2003
3.2262 | -80.2033
3.2265 | -80.2051
3.2269 | -80.208
3.2272 | -80.2097
3.2275 | -80.2124
3.2279 | -80.2145
3.2282 | -80.2177
3.2285 | -80.2194
3.2288 | -80.222
3.2292 | -80.2248
3.2295 | -80.2266
3.2298 | -80.2281
3.2302 | -80.2306
3.2305 | -80.2325
3.2308 | -80.2346
3.2311 | -80.2385
3.2315 | -80.2416
3.2318 | -80.2451
3.2321 | -80.2486
3.2325 | -80.2516
3.2328 | -80.2556
3.2331 | -80.2578
3.2335 | -80.2608
3.2338 | -80.2638
3.2341 | -80.2671
3.2344 | -80.2728
3.2348 | -80.2756
3.2351 | -80.2778
3.2354 | -80.2796
3.2358 | -80.2805
3.2361 | -80.2832
3.2364 | -80.2888
3.2367 | -80.2912
3.2371 | -80.2949
3.2374 | -80.2981
3.2377 | -80.2981
3.2381 | -80.2983
3.2384 | -80.3016
3.2387 | -80.3051
3.239 | -80.3092
3.2394 | -80.3129
3.2397 | -80.3156
3.24 | -80.3177
3.2404 | -80.3194
3.2407 | -80.3219
3.241 | -80.3239
3.2414 | -80.3283
3.2417 | -80.3311
3.242 | -80.3311
3.2423 | -80.3328
3.2427 | -80.3348
3.243 | -80.3369
3.2433 | -80.3384
3.2437 | -80.3397
3.244 | -80.3397
3.2443 | -80.3398
3.2446 | -80.3412
3.245 | -80.3429
3.2453 | -80.3447
3.2456 | -80.3479
3.246 | -80.3522
3.2463 | -80.3567
3.2466 | -80.3607
3.2469 | -80.3634
3.2473 | -80.364
3.2476 | -80.3646
3.2479 | -80.3664
3.2483 | -80.3691
3.2486 | -80.3709
3.2489 | -80.3729
3.2492 | -80.3741
3.2496 | -80.3758
3.2499 | -80.3799
3.2502 | -80.3827
3.2506 | -80.3838
3.2509 | -80.3858
3.2512 | -80.3874
3.2516 | -80.3891
3.2519 | -80.3906
3.2522 | -80.3915
3.2525 | -80.3925
3.2529 | -80.3932
3.2532 | -80.3946
3.2535 | -80.3965
3.2539 | -80.3981
3.2542 | -80.4021
3.2545 | -80.4061
3.2548 | -80.4066
3.2552 | -80.4072
3.2555 | -80.4078
3.2558 | -80.4074
3.2562 | -80.4083
3.2565 | -80.4089
3.2568 | -80.4094
3.2571 | -80.4086
3.2575 | -80.4094
3.2578 | -80.4102
3.2581 | -80.4105
3.2585 | -80.4105
3.2588 | -80.4127
3.2591 | -80.4162
3.2595 | -80.4186
3.2598 | -80.42
3.2601 | -80.4217
3.2604 | -80.4238
3.2608 | -80.4261
3.2611 | -80.4273
3.2614 | -80.4279
3.2618 | -80.4281
3.2621 | -80.4288
3.2624 | -80.4297
3.2627 | -80.4309
3.2631 | -80.4327
3.2634 | -80.4328
3.2637 | -80.4332
3.2641 | -80.4338
3.2644 | -80.4359
3.2647 | -80.4371
3.265 | -80.439
3.2654 | -80.4418
3.2657 | -80.4429
3.266 | -80.4437
3.2664 | -80.4442
3.2667 | -80.4433
3.267 | -80.443
3.2674 | -80.4432
3.2677 | -80.4434
3.268 | -80.4436
3.2683 | -80.4431
3.2687 | -80.4426
3.269 | -80.445
3.2693 | -80.449
3.2697 | -80.4543
3.27 | -80.46
3.2703 | -80.4626
3.2706 | -80.463
3.271 | -80.464
3.2713 | -80.4657
3.2716 | -80.4662
3.272 | -80.4663
3.2723 | -80.4666
3.2726 | -80.4663
3.2729 | -80.4662
3.2733 | -80.4667
3.2736 | -80.4676
3.2739 | -80.4716
3.2743 | -80.4734
3.2746 | -80.4733
3.2749 | -80.4736
3.2753 | -80.4747
3.2756 | -80.474
3.2759 | -80.4731
3.2762 | -80.4726
3.2766 | -80.4736
3.2769 | -80.4753
3.2772 | -80.4754
3.2776 | -80.4737
3.2779 | -80.4742
3.2782 | -80.4746
3.2785 | -80.4742
3.2789 | -80.4732
3.2792 | -80.472
3.2795 | -80.4714
3.2799 | -80.4703
3.2802 | -80.4684
3.2805 | -80.4671
3.2808 | -80.4654
3.2812 | -80.4642
3.2815 | -80.4638
3.2818 | -80.4648
3.2822 | -80.4654
3.2825 | -80.465
3.2828 | -80.4651
3.2832 | -80.4648
3.2835 | -80.4639
3.2838 | -80.4636
3.2841 | -80.4622
3.2845 | -80.4602
3.2848 | -80.46
3.2851 | -80.4602
3.2855 | -80.4596
3.2858 | -80.4606
3.2861 | -80.4603
3.2864 | -80.4613
3.2868 | -80.4619
3.2871 | -80.4611
3.2874 | -80.4608
3.2878 | -80.4603
3.2881 | -80.4597
3.2884 | -80.4569
3.2887 | -80.4561
3.2891 | -80.4549
3.2894 | -80.455
3.2897 | -80.456
3.2901 | -80.4564
3.2904 | -80.4555
3.2907 | -80.4556
3.291 | -80.4559
3.2914 | -80.4577
3.2917 | -80.4596
3.292 | -80.4593
3.2924 | -80.4599
3.2927 | -80.4604
3.293 | -80.4596
3.2934 | -80.4577
3.2937 | -80.4573
3.294 | -80.4566
3.2943 | -80.4568
3.2947 | -80.4572
3.295 | -80.4564
3.2953 | -80.4557
3.2957 | -80.4554
3.296 | -80.4551
3.2963 | -80.4546
3.2966 | -80.4531
3.297 | -80.449
3.2973 | -80.4473
3.2976 | -80.4469
3.298 | -80.4472
3.2983 | -80.4481
3.2986 | -80.4481
3.2989 | -80.4471
3.2993 | -80.4467
3.2996 | -80.4462
3.2999 | -80.4463
3.3003 | -80.4459
3.3006 | -80.446
3.3009 | -80.4479
3.3013 | -80.4485
3.3016 | -80.4482
3.3019 | -80.4479
3.3022 | -80.4475
3.3026 | -80.4472
3.3029 | -80.4459
3.3032 | -80.4458
3.3036 | -80.4454
3.3039 | -80.4446
3.3042 | -80.4439
3.3045 | -80.4434
3.3049 | -80.4431
3.3052 | -80.442
3.3055 | -80.4406
3.3059 | -80.4401
3.3062 | -80.4407
3.3065 | -80.4399
3.3068 | -80.4381
3.3072 | -80.4377
3.3075 | -80.4372
3.3078 | -80.4355
3.3082 | -80.4342
3.3085 | -80.4335
3.3088 | -80.4334
3.3092 | -80.4346
3.3095 | -80.4352
3.3098 | -80.4348
3.3101 | -80.4346
3.3105 | -80.4342
3.3108 | -80.4338
3.3111 | -80.4331
3.3115 | -80.4327
3.3118 | -80.4322
3.3121 | -80.4323
3.3124 | -80.4337
3.3128 | -80.4334
3.3131 | -80.433
3.3134 | -80.4337
3.3138 | -80.4338
3.3141 | -80.4344
3.3144 | -80.4346
3.3147 | -80.4352
3.3151 | -80.435
3.3154 | -80.4345
3.3157 | -80.4343
3.3161 | -80.4345
3.3164 | -80.4345
3.3167 | -80.4352
3.3171 | -80.435
3.3174 | -80.4347
3.3177 | -80.4362
3.318 | -80.4366
3.3184 | -80.4367
3.3187 | -80.4378
3.319 | -80.44
3.3194 | -80.4408
3.3197 | -80.4416
3.32 | -80.4416
3.3203 | -80.4413
3.3207 | -80.4414
3.321 | -80.4411
3.3213 | -80.4407
3.3217 | -80.4405
3.322 | -80.4414
3.3223 | -80.4428
3.3226 | -80.4431
3.323 | -80.4443
3.3233 | -80.4455
3.3236 | -80.4474
3.324 | -80.4488
3.3243 | -80.4509
3.3246 | -80.4532
3.325 | -80.4553
3.3253 | -80.456
3.3256 | -80.4551
3.3259 | -80.4555
3.3263 | -80.4556
3.3266 | -80.455
3.3269 | -80.4547
3.3273 | -80.4543
3.3276 | -80.4544
3.3279 | -80.4547
3.3282 | -80.4556
3.3286 | -80.457
3.3289 | -80.4577
3.3292 | -80.4572
3.3296 | -80.4561
3.3299 | -80.4541
3.3302 | -80.4534
3.3305 | -80.4536
3.3309 | -80.4539
3.3312 | -80.4542
3.3315 | -80.4552
3.3319 | -80.4558
3.3322 | -80.456
3.3325 | -80.4566
3.3329 | -80.4571
3.3332 | -80.4595
3.3335 | -80.4601
3.3338 | -80.4614
3.3342 | -80.4617
3.3345 | -80.4613
3.3348 | -80.4609
3.3352 | -80.4611
3.3355 | -80.4624
3.3358 | -80.4629
3.3361 | -80.4633
3.3365 | -80.4642
3.3368 | -80.4654
3.3371 | -80.4666
3.3375 | -80.4674
3.3378 | -80.4685
3.3381 | -80.4697
3.3384 | -80.4708
3.3388 | -80.472
3.3391 | -80.4718
3.3394 | -80.4721
3.3398 | -80.4737
3.3401 | -80.4745
3.3404 | -80.4754
3.3407 | -80.4762
3.3411 | -80.4772
3.3414 | -80.4777
3.3417 | -80.4793
3.3421 | -80.4815
3.3424 | -80.4827
3.3427 | -80.4858
3.3431 | -80.4863
3.3434 | -80.488
3.3437 | -80.4895
3.344 | -80.4904
3.3444 | -80.4922
3.3447 | -80.4942
3.345 | -80.4959
3.3454 | -80.4963
3.3457 | -80.4964
3.346 | -80.4967
3.3463 | -80.4995
3.3467 | -80.5029
3.347 | -80.5048
3.3473 | -80.506
3.3477 | -80.5078
3.348 | -80.5093
3.3483 | -80.5107
3.3486 | -80.5121
3.349 | -80.5128
3.3493 | -80.5146
3.3496 | -80.5157
3.35 | -80.5165
3.3503 | -80.518
3.3506 | -80.5189
3.351 | -80.5205
3.3513 | -80.5221
3.3516 | -80.5233
3.3519 | -80.5266
3.3523 | -80.528
3.3526 | -80.5279
3.3529 | -80.5304
3.3533 | -80.5324
3.3536 | -80.5352
3.3539 | -80.5363
3.3542 | -80.5357
3.3546 | -80.5341
3.3549 | -80.5331
3.3552 | -80.5338
3.3556 | -80.5361
3.3559 | -80.5379
3.3562 | -80.5392
3.3565 | -80.5405
3.3569 | -80.5422
3.3572 | -80.5439
3.3575 | -80.544
3.3579 | -80.5454
3.3582 | -80.5462
3.3585 | -80.5478
3.3589 | -80.55
3.3592 | -80.5526
3.3595 | -80.5542
3.3598 | -80.5566
3.3602 | -80.5606
3.3605 | -80.5625
3.3608 | -80.5642
3.3612 | -80.566
3.3615 | -80.5678
3.3618 | -80.5694
3.3621 | -80.5713
3.3625 | -80.5736
3.3628 | -80.5776
3.3631 | -80.5805
3.3635 | -80.5824
3.3638 | -80.5853
3.3641 | -80.587
3.3644 | -80.589
3.3648 | -80.5911
3.3651 | -80.5933
3.3654 | -80.5952
3.3658 | -80.597
3.3661 | -80.6
3.3664 | -80.6028
3.3668 | -80.6056
3.3671 | -80.6084
3.3674 | -80.6107
3.3677 | -80.6131
3.3681 | -80.6141
3.3684 | -80.6161
3.3687 | -80.619
3.3691 | -80.6218
3.3694 | -80.6241
3.3697 | -80.6256
3.37 | -80.6267
3.3704 | -80.6288
3.3707 | -80.6312
3.371 | -80.6313
3.3714 | -80.6315
3.3717 | -80.6341
3.372 | -80.637
3.3723 | -80.6405
3.3727 | -80.6432
3.373 | -80.6459
3.3733 | -80.6476
3.3737 | -80.6482
3.374 | -80.6504
3.3743 | -80.653
3.3747 | -80.6564
3.375 | -80.6592
3.3753 | -80.662
3.3756 | -80.6642
3.376 | -80.6667
3.3763 | -80.6688
3.3766 | -80.671
3.377 | -80.674
3.3773 | -80.6765
3.3776 | -80.6799
3.3779 | -80.6823
3.3783 | -80.6842
3.3786 | -80.6864
3.3789 | -80.6888
3.3793 | -80.6912
3.3796 | -80.6937
3.3799 | -80.6966
3.3802 | -80.7006
3.3806 | -80.7037
3.3809 | -80.7054
3.3812 | -80.7082
3.3816 | -80.711
3.3819 | -80.7136
3.3822 | -80.7174
3.3826 | -80.7195
3.3829 | -80.7233
3.3832 | -80.7262
3.3835 | -80.7292
3.3839 | -80.7316
3.3842 | -80.7328
3.3845 | -80.7351
3.3849 | -80.7382
3.3852 | -80.7422
3.3855 | -80.7443
3.3858 | -80.7481
3.3862 | -80.7519
3.3865 | -80.7577
3.3868 | -80.7621
3.3872 | -80.764
3.3875 | -80.7667
3.3878 | -80.7688
3.3881 | -80.7709
3.3885 | -80.7735
3.3888 | -80.7777
3.3891 | -80.781
3.3895 | -80.7831
3.3898 | -80.7855
3.3901 | -80.7886
3.3904 | -80.7917
3.3908 | -80.7945
3.3911 | -80.7998
3.3914 | -80.8052
3.3918 | -80.8083
3.3921 | -80.8101
3.3924 | -80.8116
3.3928 | -80.8148
3.3931 | -80.8175
3.3934 | -80.8219
3.3937 | -80.8244
3.3941 | -80.8275
3.3944 | -80.8285
3.3947 | -80.829
3.3951 | -80.8315
3.3954 | -80.8346
3.3957 | -80.8388
3.396 | -80.8429
3.3964 | -80.8468
3.3967 | -80.8495
3.397 | -80.8535
3.3974 | -80.8574
3.3977 | -80.8608
3.398 | -80.8647
3.3983 | -80.8672
3.3987 | -80.8697
3.399 | -80.873
3.3993 | -80.8765
3.3997 | -80.8783
3.4 | -80.8805
3.4003 | -80.8831
3.4007 | -80.8858
3.401 | -80.8884
3.4013 | -80.8912
3.4016 | -80.8935
3.402 | -80.8964
3.4023 | -80.8987
3.4026 | -80.9016
3.403 | -80.9051
3.4033 | -80.9085
3.4036 | -80.913
3.4039 | -80.9168
3.4043 | -80.9189
3.4046 | -80.9206
3.4049 | -80.9235
3.4053 | -80.9271
3.4056 | -80.9308
3.4059 | -80.9345
3.4062 | -80.9372
3.4066 | -80.9383
3.4069 | -80.94
3.4072 | -80.9422
3.4076 | -80.9446
3.4079 | -80.9454
3.4082 | -80.948
3.4086 | -80.9496
3.4089 | -80.9509
3.4092 | -80.9511
3.4095 | -80.9549
3.4099 | -80.9577
3.4102 | -80.9594
3.4105 | -80.9615
3.4109 | -80.964
3.4112 | -80.9677
3.4115 | -80.9708
3.4118 | -80.9732
3.4122 | -80.9754
3.4125 | -80.978
3.4128 | -80.9808
3.4132 | -80.9834
3.4135 | -80.9862
3.4138 | -80.9889
3.4141 | -80.991
3.4145 | -80.9934
3.4148 | -80.9963
3.4151 | -80.9997
3.4155 | -81.0038
3.4158 | -81.0077
3.4161 | -81.0099
3.4165 | -81.0126
3.4168 | -81.015
3.4171 | -81.0173
3.4174 | -81.0192
3.4178 | -81.0227
3.4181 | -81.027
3.4184 | -81.0285
3.4188 | -81.0294
3.4191 | -81.0327
3.4194 | -81.0331
3.4197 | -81.0316
3.4201 | -81.0326
3.4204 | -81.0344
3.4207 | -81.038
3.4211 | -81.0434
3.4214 | -81.0474
3.4217 | -81.0514
3.422 | -81.0543
3.4224 | -81.056
3.4227 | -81.0586
3.423 | -81.0606
3.4234 | -81.0635
3.4237 | -81.0658
3.424 | -81.0682
3.4244 | -81.0702
3.4247 | -81.0724
3.425 | -81.074
3.4253 | -81.0744
3.4257 | -81.0746
3.426 | -81.0765
3.4263 | -81.0784
3.4267 | -81.0798
3.427 | -81.0794
3.4273 | -81.0805
3.4276 | -81.0832
3.428 | -81.0851
3.4283 | -81.0868
3.4286 | -81.0887
3.429 | -81.0886
3.4293 | -81.0892
3.4296 | -81.0906
3.4299 | -81.092
3.4303 | -81.0942
3.4306 | -81.0956
3.4309 | -81.0964
3.4313 | -81.0967
3.4316 | -81.0966
3.4319 | -81.0992
3.4322 | -81.1016
3.4326 | -81.1028
3.4329 | -81.1041
3.4332 | -81.1057
3.4336 | -81.1074
3.4339 | -81.1094
3.4342 | -81.1129
3.4346 | -81.1164
3.4349 | -81.1187
3.4352 | -81.12
3.4355 | -81.1216
3.4359 | -81.1229
3.4362 | -81.1235
3.4365 | -81.1254
3.4369 | -81.1275
3.4372 | -81.1289
3.4375 | -81.1298
3.4378 | -81.132
3.4382 | -81.1348
3.4385 | -81.1366
3.4388 | -81.1384
3.4392 | -81.14
3.4395 | -81.1411
3.4398 | -81.1429
3.4402 | -81.1441
3.4405 | -81.1454
3.4408 | -81.1466
3.4411 | -81.1487
3.4415 | -81.1506
3.4418 | -81.1526
3.4421 | -81.154
3.4425 | -81.1554
3.4428 | -81.1567
3.4431 | -81.1574
3.4434 | -81.1576
3.4438 | -81.159
3.4441 | -81.1604
3.4444 | -81.1617
3.4448 | -81.163
3.4451 | -81.1643
3.4454 | -81.1663
3.4457 | -81.1671
3.4461 | -81.1686
3.4464 | -81.169
3.4467 | -81.1707
3.4471 | -81.1712
3.4474 | -81.173
3.4477 | -81.1754
3.448 | -81.1783
3.4484 | -81.1757
3.4487 | -81.1782
3.449 | -81.1824
3.4494 | -81.1822
3.4497 | -81.1821
3.45 | -81.1829
3.4504 | -81.1836
3.4507 | -81.1837
3.451 | -81.1845
3.4513 | -81.1871
3.4517 | -81.1898
3.452 | -81.1908
3.4523 | -81.1935
3.4527 | -81.1973
3.453 | -81.1991
3.4533 | -81.2011
3.4536 | -81.2035
3.454 | -81.2068
3.4543 | -81.2103
3.4546 | -81.2116
3.455 | -81.212
3.4553 | -81.2129
3.4556 | -81.2141
3.4559 | -81.216
3.4563 | -81.2179
3.4566 | -81.2171
3.4569 | -81.2191
3.4573 | -81.2195
3.4576 | -81.2194
3.4579 | -81.2188
3.4583 | -81.2183
3.4586 | -81.219
3.4589 | -81.2202
3.4592 | -81.2216
3.4596 | -81.2231
3.4599 | -81.2248
3.4602 | -81.2262
3.4606 | -81.2257
3.4609 | -81.2268
3.4612 | -81.2278
3.4615 | -81.2283
3.4619 | -81.2278
3.4622 | -81.2278
3.4625 | -81.2288
3.4629 | -81.2289
3.4632 | -81.2294
3.4635 | -81.231
3.4638 | -81.2318
3.4642 | -81.2315
3.4645 | -81.2315
3.4648 | -81.2314
3.4652 | -81.2312
3.4655 | -81.2307
3.4658 | -81.2294
3.4662 | -81.2272
3.4665 | -81.2264
3.4668 | -81.2264
3.4671 | -81.2263
3.4675 | -81.2266
3.4678 | -81.2261
3.4681 | -81.2243
3.4685 | -81.2231
3.4688 | -81.2242
3.4691 | -81.2239
3.4694 | -81.2262
3.4698 | -81.2259
3.4701 | -81.2253
3.4704 | -81.2254
3.4708 | -81.2269
3.4711 | -81.2271
3.4714 | -81.2271
3.4717 | -81.2266
3.4721 | -81.2259
3.4724 | -81.2255
3.4727 | -81.2252
3.4731 | -81.224
3.4734 | -81.2234
3.4737 | -81.2217
3.4741 | -81.2207
3.4744 | -81.2202
3.4747 | -81.2197
3.475 | -81.2195
3.4754 | -81.219
3.4757 | -81.2198
3.476 | -81.2193
3.4764 | -81.2169
3.4767 | -81.2124
3.477 | -81.2114
3.4773 | -81.2115
3.4777 | -81.213
3.478 | -81.2143
3.4783 | -81.2141
3.4787 | -81.2133
3.479 | -81.2114
3.4793 | -81.2103
3.4796 | -81.2098
3.48 | -81.2083
3.4803 | -81.2075
3.4806 | -81.2064
3.481 | -81.2064
3.4813 | -81.2066
3.4816 | -81.2059
3.4819 | -81.2048
3.4823 | -81.2014
3.4826 | -81.1999
3.4829 | -81.1972
3.4833 | -81.1958
3.4836 | -81.1947
3.4839 | -81.193
3.4843 | -81.1917
3.4846 | -81.1905
3.4849 | -81.1866
3.4852 | -81.1838
3.4856 | -81.1822
3.4859 | -81.1811
3.4862 | -81.1811
3.4866 | -81.1786
3.4869 | -81.1765
3.4872 | -81.1749
3.4875 | -81.1723
3.4879 | -81.1708
3.4882 | -81.1702
3.4885 | -81.1699
3.4889 | -81.1692
3.4892 | -81.1687
3.4895 | -81.1683
3.4898 | -81.1673
3.4902 | -81.1676
3.4905 | -81.167
3.4908 | -81.1669
3.4912 | -81.1659
3.4915 | -81.1641
3.4918 | -81.1642
3.4922 | -81.1648
3.4925 | -81.1653
3.4928 | -81.1648
3.4931 | -81.1654
3.4935 | -81.1648
3.4938 | -81.1634
3.4941 | -81.1621
3.4945 | -81.1617
3.4948 | -81.1605
3.4951 | -81.1588
3.4954 | -81.1579
3.4958 | -81.1569
3.4961 | -81.1558
3.4964 | -81.1552
3.4968 | -81.154
3.4971 | -81.1519
3.4974 | -81.1506
3.4977 | -81.1497
3.4981 | -81.1486
3.4984 | -81.148
3.4987 | -81.1471
3.4991 | -81.146
3.4994 | -81.1435
3.4997 | -81.141
3.5001 | -81.1385
3.5004 | -81.1359
3.5007 | -81.1341
3.501 | -81.1326
3.5014 | -81.1314
3.5017 | -81.1296
3.502 | -81.1285
3.5024 | -81.1285
3.5027 | -81.1268
3.503 | -81.1256
3.5033 | -81.1242
3.5037 | -81.1231
3.504 | -81.1218
3.5043 | -81.1219
3.5047 | -81.1202
3.505 | -81.1184
3.5053 | -81.1174
3.5056 | -81.1156
3.506 | -81.1134
3.5063 | -81.1111
3.5066 | -81.1098
3.507 | -81.1078
3.5073 | -81.1062
3.5076 | -81.1048
3.508 | -81.1032
3.5083 | -81.1013
3.5086 | -81.0985
3.5089 | -81.0968
3.5093 | -81.0956
3.5096 | -81.0938
3.5099 | -81.0914
3.5103 | -81.0896
3.5106 | -81.0883
3.5109 | -81.0875
3.5112 | -81.0868
3.5116 | -81.0858
3.5119 | -81.0856
3.5122 | -81.0835
3.5126 | -81.0808
3.5129 | -81.0786
3.5132 | -81.0781
3.5135 | -81.0753
3.5139 | -81.0765
3.5142 | -81.0753
3.5145 | -81.0735
3.5149 | -81.0715
3.5152 | -81.0703
3.5155 | -81.0692
3.5159 | -81.0684
3.5162 | -81.0681
3.5165 | -81.0681
3.5168 | -81.0678
3.5172 | -81.0677
3.5175 | -81.0666
3.5178 | -81.0654
3.5182 | -81.0639
3.5185 | -81.0625
3.5188 | -81.0593
3.5191 | -81.0558
3.5195 | -81.054
3.5198 | -81.0522
3.5201 | -81.0499
3.5205 | -81.049
3.5208 | -81.0486
3.5211 | -81.0459
3.5214 | -81.0445
3.5218 | -81.0435
3.5221 | -81.041
3.5224 | -81.0387
3.5228 | -81.0371
3.5231 | -81.0361
3.5234 | -81.0344
3.5238 | -81.0329
3.5241 | -81.0323
3.5244 | -81.033
3.5247 | -81.032
3.5251 | -81.031
3.5254 | -81.0307
3.5257 | -81.0277
3.5261 | -81.0267
3.5264 | -81.025
3.5267 | -81.0224
3.527 | -81.0199
3.5274 | -81.018
3.5277 | -81.0157
3.528 | -81.0109
3.5284 | -81.0063
3.5287 | -81.0041
3.529 | -81.0029
3.5293 | -81.0008
3.5297 | -80.9985
3.53 | -80.9965
3.5303 | -80.9955
3.5307 | -80.9935
3.531 | -80.9922
3.5313 | -80.9904
3.5316 | -80.9884
3.532 | -80.9859
3.5323 | -80.9832
3.5326 | -80.9803
3.533 | -80.9786
3.5333 | -80.9756
3.5336 | -80.973
3.534 | -80.971
3.5343 | -80.9695
3.5346 | -80.972
3.5349 | -80.9712
3.5353 | -80.9681
3.5356 | -80.9655
3.5359 | -80.9614
3.5363 | -80.9574
3.5366 | -80.9553
3.5369 | -80.9533
3.5372 | -80.9514
3.5376 | -80.9494
3.5379 | -80.947
3.5382 | -80.9445
3.5386 | -80.9427
3.5389 | -80.9415
3.5392 | -80.9397
3.5395 | -80.9379
3.5399 | -80.9337
3.5402 | -80.9316
3.5405 | -80.9301
3.5409 | -80.9282
3.5412 | -80.9266
3.5415 | -80.9251
3.5419 | -80.924
3.5422 | -80.922
3.5425 | -80.9206
3.5428 | -80.9189
3.5432 | -80.9172
3.5435 | -80.9166
3.5438 | -80.9158
3.5442 | -80.9134
3.5445 | -80.9111
3.5448 | -80.9104
3.5451 | -80.9099
3.5455 | -80.9079
3.5458 | -80.9066
3.5461 | -80.9042
3.5465 | -80.9026
3.5468 | -80.9008
3.5471 | -80.8981
3.5474 | -80.8961
3.5478 | -80.8936
3.5481 | -80.8922
3.5484 | -80.8908
3.5488 | -80.8888
3.5491 | -80.8865
3.5494 | -80.8848
3.5498 | -80.883
3.5501 | -80.881
3.5504 | -80.8795
3.5507 | -80.8768
3.5511 | -80.8746
3.5514 | -80.8715
3.5517 | -80.8691
3.5521 | -80.8673
3.5524 | -80.866
3.5527 | -80.8647
3.553 | -80.863
3.5534 | -80.8614
3.5537 | -80.8601
3.554 | -80.8586
3.5544 | -80.8562
3.5547 | -80.8546
3.555 | -80.8529
3.5553 | -80.8521
3.5557 | -80.8506
3.556 | -80.8491
3.5563 | -80.8467
3.5567 | -80.8431
3.557 | -80.842
3.5573 | -80.8394
3.5577 | -80.834
3.558 | -80.8317
3.5583 | -80.8292
3.5586 | -80.8259
3.559 | -80.8228
3.5593 | -80.8207
3.5596 | -80.819
3.56 | -80.8182
3.5603 | -80.818
3.5606 | -80.8169
3.5609 | -80.8163
3.5613 | -80.8146
3.5616 | -80.813
3.5619 | -80.8111
3.5623 | -80.8101
3.5626 | -80.8086
3.5629 | -80.8067
3.5632 | -80.8052
3.5636 | -80.8035
3.5639 | -80.8008
3.5642 | -80.7972
3.5646 | -80.7954
3.5649 | -80.7932
3.5652 | -80.792
3.5656 | -80.7893
3.5659 | -80.7871
3.5662 | -80.7867
3.5665 | -80.7839
3.5669 | -80.782
3.5672 | -80.7805
3.5675 | -80.7788
3.5679 | -80.777
3.5682 | -80.7754
3.5685 | -80.7738
3.5688 | -80.7726
3.5692 | -80.7732
3.5695 | -80.774
3.5698 | -80.7729
3.5702 | -80.7715
3.5705 | -80.7676
3.5708 | -80.7646
3.5711 | -80.7628
3.5715 | -80.7604
3.5718 | -80.7573
3.5721 | -80.7551
3.5725 | -80.7532
3.5728 | -80.7518
3.5731 | -80.7502
3.5734 | -80.7486
3.5738 | -80.7475
3.5741 | -80.7462
3.5744 | -80.7451
3.5748 | -80.7439
3.5751 | -80.7431
3.5754 | -80.7414
3.5758 | -80.7403
3.5761 | -80.7395
3.5764 | -80.7384
3.5767 | -80.7367
3.5771 | -80.7344
3.5774 | -80.7326
3.5777 | -80.7313
3.5781 | -80.7296
3.5784 | -80.7274
3.5787 | -80.7258
3.579 | -80.7241
3.5794 | -80.7234
3.5797 | -80.7226
3.58 | -80.723
3.5804 | -80.7218
3.5807 | -80.7205
3.581 | -80.7178
3.5814 | -80.7163
3.5817 | -80.7151
3.582 | -80.7142
3.5823 | -80.7125
3.5827 | -80.7116
3.583 | -80.7103
3.5833 | -80.7089
3.5837 | -80.7068
3.584 | -80.7058
3.5843 | -80.7053
3.5846 | -80.7041
3.585 | -80.7035
3.5853 | -80.7032
3.5856 | -80.7025
3.586 | -80.7019
3.5863 | -80.701
3.5866 | -80.7009
3.5869 | -80.7008
3.5873 | -80.7001
3.5876 | -80.6993
3.5879 | -80.6985
3.5883 | -80.6976
3.5886 | -80.697
3.5889 | -80.6963
3.5892 | -80.6941
3.5896 | -80.6943
3.5899 | -80.6929
3.5902 | -80.6918
3.5906 | -80.6911
3.5909 | -80.6909
3.5912 | -80.6908
3.5916 | -80.6909
3.5919 | -80.6909
3.5922 | -80.6903
3.5925 | -80.6884
3.5929 | -80.6867
3.5932 | -80.6854
3.5935 | -80.6847
3.5939 | -80.6844
3.5942 | -80.6851
3.5945 | -80.6843
3.5948 | -80.6854
3.5952 | -80.6852
3.5955 | -80.6847
3.5958 | -80.6857
3.5962 | -80.6879
3.5965 | -80.6892
3.5968 | -80.6894
3.5971 | -80.6885
3.5975 | -80.6877
3.5978 | -80.687
3.5981 | -80.6862
3.5985 | -80.6849
3.5988 | -80.6835
3.5991 | -80.6833
3.5995 | -80.6851
3.5998 | -80.6849
3.6001 | -80.684
3.6004 | -80.6816
3.6008 | -80.6805
3.6011 | -80.6801
3.6014 | -80.6804
3.6018 | -80.681
3.6021 | -80.6792
3.6024 | -80.6783
3.6027 | -80.6769
3.6031 | -80.677
3.6034 | -80.6755
3.6037 | -80.6765
3.6041 | -80.6764
3.6044 | -80.676
3.6047 | -80.6755
3.605 | -80.6751
3.6054 | -80.6748
3.6057 | -80.6742
3.606 | -80.6735
3.6064 | -80.6735
3.6067 | -80.6733
3.607 | -80.6728
3.6074 | -80.6722
3.6077 | -80.6718
3.608 | -80.6716
3.6083 | -80.6728
3.6087 | -80.6728
3.609 | -80.6725
3.6093 | -80.6726
3.6097 | -80.6726
3.61 | -80.6721
3.6103 | -80.6719
3.6106 | -80.6727
3.611 | -80.6728
3.6113 | -80.6727
3.6116 | -80.6728
3.612 | -80.6725
3.6123 | -80.6731
3.6126 | -80.6747
3.6129 | -80.6749
3.6133 | -80.6751
3.6136 | -80.6753
3.6139 | -80.6783
3.6143 | -80.6796
3.6146 | -80.6781
3.6149 | -80.6783
3.6153 | -80.6782
3.6156 | -80.6781
3.6159 | -80.6777
3.6162 | -80.6781
3.6166 | -80.6801
3.6169 | -80.681
3.6172 | -80.6803
3.6176 | -80.6806
3.6179 | -80.6821
3.6182 | -80.6824
3.6185 | -80.6827
3.6189 | -80.6821
3.6192 | -80.6828
3.6195 | -80.6836
3.6199 | -80.684
3.6202 | -80.6842
3.6205 | -80.6845
3.6208 | -80.685
3.6212 | -80.6858
3.6215 | -80.6865
3.6218 | -80.6875
3.6222 | -80.6879
3.6225 | -80.6876
3.6228 | -80.6874
3.6231 | -80.689
3.6235 | -80.6904
3.6238 | -80.6914
3.6241 | -80.6929
3.6245 | -80.6941
3.6248 | -80.695
3.6251 | -80.6949
3.6255 | -80.6954
3.6258 | -80.696
3.6261 | -80.6969
3.6264 | -80.6973
3.6268 | -80.6984
3.6271 | -80.6992
3.6274 | -80.7017
3.6278 | -80.7029
3.6281 | -80.7055
3.6284 | -80.7066
3.6287 | -80.7069
3.6291 | -80.708
3.6294 | -80.709
3.6297 | -80.7096
3.6301 | -80.7104
3.6304 | -80.7116
3.6307 | -80.7129
3.6311 | -80.7142
3.6314 | -80.7144
3.6317 | -80.715
3.632 | -80.7158
3.6324 | -80.7161
3.6327 | -80.7169
3.633 | -80.7185
3.6334 | -80.7206
3.6337 | -80.7223
3.634 | -80.7232
3.6343 | -80.7248
3.6347 | -80.7262
3.635 | -80.7273
3.6353 | -80.7286
3.6357 | -80.7296
3.636 | -80.7305
3.6363 | -80.7322
3.6366 | -80.7331
3.637 | -80.7343
3.6373 | -80.7353
3.6376 | -80.7365
3.638 | -80.7361
3.6383 | -80.7361
3.6386 | -80.7368
3.6389 | -80.7392
3.6393 | -80.7397
3.6396 | -80.7402
3.6399 | -80.7414
3.6403 | -80.7427
3.6406 | -80.7438
3.6409 | -80.7446
3.6413 | -80.7462
3.6416 | -80.7474
3.6419 | -80.7491
3.6422 | -80.7506
3.6426 | -80.7518
3.6429 | -80.7533
3.6432 | -80.7549
3.6436 | -80.7561
3.6439 | -80.7573
3.6442 | -80.7584
3.6445 | -80.7606
3.6449 | -80.7607
3.6452 | -80.762
3.6455 | -80.7629
3.6459 | -80.7642
3.6462 | -80.7655
3.6465 | -80.7685
3.6468 | -80.7706
3.6472 | -80.7727
3.6475 | -80.7746
3.6478 | -80.7765
3.6482 | -80.7781
3.6485 | -80.7793
3.6488 | -80.7792
3.6492 | -80.7793
3.6495 | -80.7814
3.6498 | -80.7837
3.6501 | -80.7847
3.6505 | -80.7844
3.6508 | -80.7852
3.6511 | -80.7865
3.6515 | -80.7875
3.6518 | -80.7903
3.6521 | -80.7926
3.6524 | -80.794
3.6528 | -80.7968
3.6531 | -80.801
3.6534 | -80.8027
3.6538 | -80.8051
3.6541 | -80.8075
3.6544 | -80.8098
3.6547 | -80.8119
3.6551 | -80.8132
3.6554 | -80.8152
3.6557 | -80.8176
3.6561 | -80.8204
3.6564 | -80.8223
3.6567 | -80.8236
3.6571 | -80.8254
3.6574 | -80.8271
3.6577 | -80.8295
3.658 | -80.8323
3.6584 | -80.8338
3.6587 | -80.8352
3.659 | -80.8366
3.6594 | -80.8386
3.6597 | -80.8403
3.66 | -80.8418
3.6603 | -80.842
3.6607 | -80.843
3.661 | -80.8441
3.6613 | -80.8459
3.6617 | -80.8478
3.662 | -80.8494
3.6623 | -80.8507
3.6626 | -80.8527
3.663 | -80.8548
3.6633 | -80.8579
3.6636 | -80.8605
3.664 | -80.8625
3.6643 | -80.8646
3.6646 | -80.8675
3.665 | -80.8707
3.6653 | -80.8725
3.6656 | -80.8738
3.6659 | -80.8774
3.6663 | -80.883
3.6666 | -80.886
3.6669 | -80.8866
3.6673 | -80.8866
3.6676 | -80.8883
3.6679 | -80.89
3.6682 | -80.8888
3.6686 | -80.8887
3.6689 | -80.891
3.6692 | -80.8921
3.6696 | -80.8921
3.6699 | -80.8939
3.6702 | -80.8944
3.6705 | -80.8952
3.6709 | -80.8969
3.6712 | -80.8992
3.6715 | -80.9021
3.6719 | -80.9059
3.6722 | -80.9088
3.6725 | -80.9097
3.6728 | -80.9114
3.6732 | -80.9135
3.6735 | -80.9173
3.6738 | -80.9195
3.6742 | -80.9216
3.6745 | -80.9233
3.6748 | -80.9262
3.6752 | -80.928
3.6755 | -80.9299
3.6758 | -80.932
3.6761 | -80.9356
3.6765 | -80.9393
3.6768 | -80.9408
3.6771 | -80.9429
3.6775 | -80.9448
3.6778 | -80.9476
3.6781 | -80.95
3.6784 | -80.9534
3.6788 | -80.9586
3.6791 | -80.961
3.6794 | -80.9623
3.6798 | -80.9644
3.6801 | -80.9667
3.6804 | -80.9686
3.6807 | -80.9701
3.6811 | -80.9712
3.6814 | -80.9728
3.6817 | -80.9721
3.6821 | -80.9739
3.6824 | -80.975
3.6827 | -80.9753
3.6831 | -80.9772
3.6834 | -80.9797
3.6837 | -80.9812
3.684 | -80.983
3.6844 | -80.985
3.6847 | -80.9862
3.685 | -80.9874
3.6854 | -80.989
3.6857 | -80.9908
3.686 | -80.9925
3.6863 | -80.9945
3.6867 | -80.9962
3.687 | -80.9969
3.6873 | -80.9964
3.6877 | -80.9973
3.688 | -80.9991
3.6883 | -81.001
3.6886 | -81.0032
3.689 | -81.0052
3.6893 | -81.0063
3.6896 | -81.0074
3.69 | -81.0085
3.6903 | -81.0103
3.6906 | -81.0119
3.691 | -81.0137
3.6913 | -81.0162
3.6916 | -81.018
3.6919 | -81.0196
3.6923 | -81.0204
3.6926 | -81.022
3.6929 | -81.024
3.6933 | -81.0259
3.6936 | -81.0288
3.6939 | -81.0303
3.6942 | -81.031
3.6946 | -81.0318
3.6949 | -81.0337
3.6952 | -81.0358
3.6956 | -81.0382
3.6959 | -81.0401
3.6962 | -81.0414
3.6965 | -81.0423
3.6969 | -81.0478
3.6972 | -81.0511
3.6975 | -81.0534
3.6979 | -81.0566
3.6982 | -81.0594
3.6985 | -81.0613
3.6989 | -81.0625
3.6992 | -81.0636
3.6995 | -81.0654
3.6998 | -81.0674
3.7002 | -81.0693
3.7005 | -81.0692
3.7008 | -81.0703
3.7012 | -81.071
3.7015 | -81.0725
3.7018 | -81.0741
3.7021 | -81.0761
3.7025 | -81.0781
3.7028 | -81.0795
3.7031 | -81.0807
3.7035 | -81.0773
3.7038 | -81.0782
3.7041 | -81.0798
3.7044 | -81.0817
3.7048 | -81.0835
3.7051 | -81.0851
3.7054 | -81.0868
3.7058 | -81.0892
3.7061 | -81.092
3.7064 | -81.0929
3.7068 | -81.0931
3.7071 | -81.0952
3.7074 | -81.0975
3.7077 | -81.0989
3.7081 | -81.1002
3.7084 | -81.1004
3.7087 | -81.1002
3.7091 | -81.1
3.7094 | -81.1015
3.7097 | -81.1035
3.71 | -81.1046
3.7104 | -81.1055
3.7107 | -81.1046
3.711 | -81.1035
3.7114 | -81.105
3.7117 | -81.1067
3.712 | -81.1076
3.7123 | -81.1089
3.7127 | -81.1095
3.713 | -81.111
3.7133 | -81.1123
3.7137 | -81.1144
3.714 | -81.1157
3.7143 | -81.1166
3.7147 | -81.1181
3.715 | -81.1208
3.7153 | -81.1233
3.7156 | -81.1259
3.716 | -81.1283
3.7163 | -81.1306
3.7166 | -81.1317
3.717 | -81.1332
3.7173 | -81.1345
3.7176 | -81.1357
3.7179 | -81.1385
3.7183 | -81.1415
3.7186 | -81.1436
3.7189 | -81.1471
3.7193 | -81.1479
3.7196 | -81.1496
3.7199 | -81.1508
3.7202 | -81.1513
3.7206 | -81.1528
3.7209 | -81.1551
3.7212 | -81.1561
3.7216 | -81.1576
3.7219 | -81.1593
3.7222 | -81.1606
3.7226 | -81.1588
3.7229 | -81.159
3.7232 | -81.1597
3.7235 | -81.1601
3.7239 | -81.1626
3.7242 | -81.1643
3.7245 | -81.1653
3.7249 | -81.1655
3.7252 | -81.1661
3.7255 | -81.1658
3.7258 | -81.1667
3.7262 | -81.1671
3.7265 | -81.1662
3.7268 | -81.1678
3.7272 | -81.1724
3.7275 | -81.174
3.7278 | -81.1743
3.7281 | -81.1732
3.7285 | -81.1733
3.7288 | -81.1741
3.7291 | -81.1752
3.7295 | -81.1769
3.7298 | -81.1782
3.7301 | -81.1783
3.7304 | -81.1797
3.7308 | -81.181
3.7311 | -81.1826
3.7314 | -81.1842
3.7318 | -81.1848
3.7321 | -81.1822
3.7324 | -81.1826
3.7328 | -81.1843
3.7331 | -81.1878
3.7334 | -81.1888
3.7337 | -81.1911
3.7341 | -81.192
3.7344 | -81.191
3.7347 | -81.1905
3.7351 | -81.1915
3.7354 | -81.1922
3.7357 | -81.1933
3.736 | -81.196
3.7364 | -81.1981
3.7367 | -81.1992
3.737 | -81.1992
3.7374 | -81.1995
3.7377 | -81.2024
3.738 | -81.2044
3.7383 | -81.2054
3.7387 | -81.2074
3.739 | -81.2083
3.7393 | -81.2088
3.7397 | -81.2088
3.74 | -81.2087
3.7403 | -81.2092
3.7407 | -81.2098
3.741 | -81.2112
3.7413 | -81.2103
3.7416 | -81.2076
3.742 | -81.2057
3.7423 | -81.2043
3.7426 | -81.205
3.743 | -81.2063
3.7433 | -81.2063
3.7436 | -81.2065
3.7439 | -81.206
3.7443 | -81.2061
3.7446 | -81.2063
3.7449 | -81.2063
3.7453 | -81.2053
3.7456 | -81.2047
3.7459 | -81.2048
3.7462 | -81.2049
3.7466 | -81.2048
3.7469 | -81.205
3.7472 | -81.2044
3.7476 | -81.2044
3.7479 | -81.205
3.7482 | -81.2047
3.7486 | -81.2055
3.7489 | -81.2068
3.7492 | -81.2071
3.7495 | -81.2073
3.7499 | -81.2077
3.7502 | -81.2084
3.7505 | -81.2083
3.7509 | -81.208
3.7512 | -81.2064
3.7515 | -81.2074
3.7518 | -81.2073
3.7522 | -81.2059
3.7525 | -81.2055
3.7528 | -81.2055
3.7532 | -81.2054
3.7535 | -81.2045
3.7538 | -81.2049
3.7541 | -81.2059
3.7545 | -81.2061
3.7548 | -81.2065
3.7551 | -81.2064
3.7555 | -81.2072
3.7558 | -81.2084
3.7561 | -81.2088
3.7565 | -81.209
3.7568 | -81.2089
3.7571 | -81.2094
3.7574 | -81.2105
3.7578 | -81.2108
3.7581 | -81.2114
3.7584 | -81.2118
3.7588 | -81.2111
3.7591 | -81.2097
3.7594 | -81.2109
3.7597 | -81.2108
3.7601 | -81.2101
3.7604 | -81.2109
3.7607 | -81.2103
3.7611 | -81.2106
3.7614 | -81.2093
3.7617 | -81.2053
3.762 | -81.204
3.7624 | -81.2038
3.7627 | -81.2043
3.763 | -81.2045
3.7634 | -81.2039
3.7637 | -81.2018
3.764 | -81.201
3.7643 | -81.201
3.7647 | -81.1995
3.765 | -81.2001
3.7653 | -81.2002
3.7657 | -81.2006
3.766 | -81.2012
3.7663 | -81.2013
3.7667 | -81.2012
3.767 | -81.2009
3.7673 | -81.2006
3.7676 | -81.1997
3.768 | -81.199
3.7683 | -81.1983
3.7686 | -81.1989
3.769 | -81.1997
3.7693 | -81.1995
3.7696 | -81.198
3.7699 | -81.1969
3.7703 | -81.196
3.7706 | -81.1948
3.7709 | -81.1945
3.7713 | -81.194
3.7716 | -81.1934
3.7719 | -81.1928
3.7723 | -81.1939
3.7726 | -81.1951
3.7729 | -81.195
3.7732 | -81.1951
3.7736 | -81.1929
3.7739 | -81.1911
3.7742 | -81.1915
3.7746 | -81.1921
3.7749 | -81.1907
3.7752 | -81.1893
3.7755 | -81.1896
3.7759 | -81.1895
3.7762 | -81.189
3.7765 | -81.1881
3.7769 | -81.1878
3.7772 | -81.1862
3.7775 | -81.1843
3.7778 | -81.183
3.7782 | -81.182
3.7785 | -81.1812
3.7788 | -81.1804
3.7792 | -81.1806
3.7795 | -81.1791
3.7798 | -81.1799
3.7801 | -81.1802
3.7805 | -81.1805
3.7808 | -81.1807
3.7811 | -81.1806
3.7815 | -81.1808
3.7818 | -81.181
3.7821 | -81.1833
3.7825 | -81.1854
3.7828 | -81.1844
3.7831 | -81.1827
3.7834 | -81.1808
3.7838 | -81.1798
3.7841 | -81.1795
3.7844 | -81.1793
3.7848 | -81.1787
3.7851 | -81.178
3.7854 | -81.1773
3.7857 | -81.1766
3.7861 | -81.1768
3.7864 | -81.1775
3.7867 | -81.1773
3.7871 | -81.178
3.7874 | -81.1773
3.7877 | -81.1779
3.788 | -81.1774
3.7884 | -81.1773
3.7887 | -81.1764
3.789 | -81.1743
3.7894 | -81.1718
3.7897 | -81.1694
3.79 | -81.1697
3.7904 | -81.1695
3.7907 | -81.1693
3.791 | -81.1674
3.7913 | -81.1667
3.7917 | -81.1656
3.792 | -81.1648
3.7923 | -81.164
3.7927 | -81.1632
3.793 | -81.1629
3.7933 | -81.163
3.7936 | -81.1624
3.794 | -81.1612
3.7943 | -81.1577
3.7946 | -81.158
3.795 | -81.1572
3.7953 | -81.1531
3.7956 | -81.1487
3.7959 | -81.1473
3.7963 | -81.1457
3.7966 | -81.1444
3.7969 | -81.1436
3.7973 | -81.1431
3.7976 | -81.1425
3.7979 | -81.1417
3.7983 | -81.1409
3.7986 | -81.14
3.7989 | -81.1384
3.7992 | -81.1373
3.7996 | -81.1368
3.7999 | -81.1358
3.8002 | -81.1332
3.8006 | -81.1305
3.8009 | -81.1293
3.8012 | -81.1285
3.8015 | -81.1302
3.8019 | -81.1306
3.8022 | -81.13
3.8025 | -81.1296
3.8029 | -81.1281
3.8032 | -81.127
3.8035 | -81.1261
3.8038 | -81.1252
3.8042 | -81.1233
3.8045 | -81.1211
3.8048 | -81.1199
3.8052 | -81.118
3.8055 | -81.1163
3.8058 | -81.1152
3.8062 | -81.1146
3.8065 | -81.1127
3.8068 | -81.1125
3.8071 | -81.1127
3.8075 | -81.1121
3.8078 | -81.1096
3.8081 | -81.1066
3.8085 | -81.1052
3.8088 | -81.1045
3.8091 | -81.1034
3.8094 | -81.1021
3.8098 | -81.101
3.8101 | -81.0996
3.8104 | -81.0983
3.8108 | -81.0965
3.8111 | -81.0944
3.8114 | -81.0926
3.8117 | -81.0914
3.8121 | -81.0896
3.8124 | -81.0895
3.8127 | -81.0905
3.8131 | -81.0894
3.8134 | -81.0894
3.8137 | -81.0919
3.814 | -81.091
3.8144 | -81.0897
3.8147 | -81.09
3.815 | -81.0903
3.8154 | -81.0898
3.8157 | -81.0887
3.816 | -81.0878
3.8164 | -81.0858
3.8167 | -81.0823
3.817 | -81.0801
3.8173 | -81.079
3.8177 | -81.078
3.818 | -81.0764
3.8183 | -81.0739
3.8187 | -81.0716
3.819 | -81.0704
3.8193 | -81.0695
3.8196 | -81.0683
3.82 | -81.0669
3.8203 | -81.0657
3.8206 | -81.0655
3.821 | -81.0664
3.8213 | -81.0652
3.8216 | -81.0649
3.822 | -81.0638
3.8223 | -81.0622
3.8226 | -81.0628
3.8229 | -81.062
3.8233 | -81.0629
3.8236 | -81.063
3.8239 | -81.0618
3.8243 | -81.0608
3.8246 | -81.0588
3.8249 | -81.0565
3.8252 | -81.0555
3.8256 | -81.0548
3.8259 | -81.0535
3.8262 | -81.0531
3.8266 | -81.0528
3.8269 | -81.0521
3.8272 | -81.0518
3.8275 | -81.0525
3.8279 | -81.0528
3.8282 | -81.0515
3.8285 | -81.05
3.8289 | -81.0492
3.8292 | -81.0495
3.8295 | -81.0482
3.8298 | -81.0468
3.8302 | -81.0455
3.8305 | -81.0445
3.8308 | -81.0435
3.8312 | -81.0419
3.8315 | -81.0409
3.8318 | -81.038
3.8322 | -81.0373
3.8325 | -81.0368
3.8328 | -81.0363
3.8331 | -81.0352
3.8335 | -81.0318
3.8338 | -81.0276
3.8341 | -81.026
3.8345 | -81.0264
3.8348 | -81.0258
3.8351 | -81.0248
3.8354 | -81.0235
3.8358 | -81.0228
3.8361 | -81.022
3.8364 | -81.0214
3.8368 | -81.0205
3.8371 | -81.0192
3.8374 | -81.0181
3.8377 | -81.0166
3.8381 | -81.0147
3.8384 | -81.0132
3.8387 | -81.012
3.8391 | -81.0113
3.8394 | -81.0098
3.8397 | -81.0087
3.8401 | -81.0078
3.8404 | -81.0071
3.8407 | -81.0056
3.841 | -81.0044
3.8414 | -81.0035
3.8417 | -81.0035
3.842 | -81.0035
3.8424 | -81.0021
3.8427 | -81.001
3.843 | -80.9998
3.8433 | -80.9983
3.8437 | -80.9973
3.844 | -80.9962
3.8443 | -80.9957
3.8447 | -80.9949
3.845 | -80.9939
3.8453 | -80.993
3.8456 | -80.9911
3.846 | -80.9851
3.8463 | -80.9841
3.8466 | -80.9833
3.847 | -80.9823
3.8473 | -80.9806
3.8476 | -80.9804
3.848 | -80.9802
3.8483 | -80.9795
3.8486 | -80.9793
3.8489 | -80.9792
3.8493 | -80.9771
3.8496 | -80.9754
3.8499 | -80.9747
3.8503 | -80.9743
3.8506 | -80.9732
3.8509 | -80.9712
3.8512 | -80.9716
3.8516 | -80.9724
3.8519 | -80.9723
3.8522 | -80.9716
3.8526 | -80.9708
3.8529 | -80.9703
3.8532 | -80.9697
3.8535 | -80.9688
3.8539 | -80.9677
3.8542 | -80.9678
3.8545 | -80.9664
3.8549 | -80.965
3.8552 | -80.9646
3.8555 | -80.9629
3.8559 | -80.9609
3.8562 | -80.9607
3.8565 | -80.9599
3.8568 | -80.9586
3.8572 | -80.9578
3.8575 | -80.958
3.8578 | -80.9586
3.8582 | -80.959
3.8585 | -80.9578
3.8588 | -80.956
3.8591 | -80.9548
3.8595 | -80.9551
3.8598 | -80.9564
3.8601 | -80.9563
3.8605 | -80.9551
3.8608 | -80.9539
3.8611 | -80.953
3.8614 | -80.9528
3.8618 | -80.9516
3.8621 | -80.952
3.8624 | -80.951
3.8628 | -80.9513
3.8631 | -80.9506
3.8634 | -80.9497
3.8638 | -80.9488
3.8641 | -80.9474
3.8644 | -80.9454
3.8647 | -80.9442
3.8651 | -80.9411
3.8654 | -80.9387
3.8657 | -80.9383
3.8661 | -80.937
3.8664 | -80.9362
3.8667 | -80.9348
3.867 | -80.9324
3.8674 | -80.9311
3.8677 | -80.93
3.868 | -80.9292
3.8684 | -80.9283
3.8687 | -80.926
3.869 | -80.9235
3.8693 | -80.9181
3.8697 | -80.9162
3.87 | -80.9155
3.8703 | -80.9136
3.8707 | -80.9127
3.871 | -80.9119
3.8713 | -80.9099
3.8716 | -80.9083
3.872 | -80.9078
3.8723 | -80.9071
3.8726 | -80.9049
3.873 | -80.9032
3.8733 | -80.9019
3.8736 | -80.9012
3.874 | -80.9005
3.8743 | -80.9002
3.8746 | -80.899
3.8749 | -80.8993
3.8753 | -80.8987
3.8756 | -80.8979
3.8759 | -80.8966
3.8763 | -80.8959
3.8766 | -80.8957
3.8769 | -80.8955
3.8772 | -80.8953
3.8776 | -80.8948
3.8779 | -80.8936
3.8782 | -80.8917
3.8786 | -80.8912
3.8789 | -80.8907
3.8792 | -80.8898
3.8795 | -80.8899
3.8799 | -80.8887
3.8802 | -80.8871
3.8805 | -80.8868
3.8809 | -80.8855
3.8812 | -80.8846
3.8815 | -80.884
3.8819 | -80.8836
3.8822 | -80.8835
3.8825 | -80.8832
3.8828 | -80.8821
3.8832 | -80.8818
3.8835 | -80.8815
3.8838 | -80.881
3.8842 | -80.8804
3.8845 | -80.8791
3.8848 | -80.8783
3.8851 | -80.8773
3.8855 | -80.8769
3.8858 | -80.8765
3.8861 | -80.876
3.8865 | -80.8745
3.8868 | -80.872
3.8871 | -80.8695
3.8874 | -80.8683
3.8878 | -80.8672
3.8881 | -80.8662
3.8884 | -80.8653
3.8888 | -80.8634
3.8891 | -80.8613
3.8894 | -80.86
3.8898 | -80.8589
3.8901 | -80.858
3.8904 | -80.8573
3.8907 | -80.8573
3.8911 | -80.8565
3.8914 | -80.856
3.8917 | -80.8548
3.8921 | -80.8541
3.8924 | -80.8535
3.8927 | -80.8533
3.893 | -80.8536
3.8934 | -80.8525
3.8937 | -80.8529
3.894 | -80.8532
3.8944 | -80.8532
3.8947 | -80.8528
3.895 | -80.8516
3.8953 | -80.8505
3.8957 | -80.8502
3.896 | -80.8491
3.8963 | -80.8473
3.8967 | -80.846
3.897 | -80.8449
3.8973 | -80.8443
3.8977 | -80.8433
3.898 | -80.8425
3.8983 | -80.8437
3.8986 | -80.8446
3.899 | -80.8436
3.8993 | -80.8429
3.8996 | -80.842
3.9 | -80.8433
3.9003 | -80.8444
3.9006 | -80.8447
3.9009 | -80.8452
3.9013 | -80.8464
3.9016 | -80.8453
3.9019 | -80.8438
3.9023 | -80.8431
3.9026 | -80.8428
3.9029 | -80.8431
3.9032 | -80.8428
3.9036 | -80.841
3.9039 | -80.8405
3.9042 | -80.8397
3.9046 | -80.8391
3.9049 | -80.8383
3.9052 | -80.8374
3.9056 | -80.8369
3.9059 | -80.8366
3.9062 | -80.8362
3.9065 | -80.8358
3.9069 | -80.8356
3.9072 | -80.8365
3.9075 | -80.836
3.9079 | -80.8362
3.9082 | -80.8361
3.9085 | -80.8357
3.9088 | -80.8355
3.9092 | -80.8352
3.9095 | -80.8344
3.9098 | -80.8341
3.9102 | -80.8336
3.9105 | -80.8316
3.9108 | -80.8299
3.9111 | -80.8298
3.9115 | -80.8308
3.9118 | -80.8317
3.9121 | -80.8313
3.9125 | -80.8317
3.9128 | -80.8317
3.9131 | -80.8323
3.9135 | -80.8329
3.9138 | -80.8332
3.9141 | -80.8332
3.9144 | -80.8325
3.9148 | -80.8325
3.9151 | -80.8324
3.9154 | -80.8325
3.9158 | -80.8312
3.9161 | -80.8306
3.9164 | -80.8292
3.9167 | -80.8284
3.9171 | -80.8285
3.9174 | -80.8289
3.9177 | -80.8297
3.9181 | -80.8294
3.9184 | -80.828
3.9187 | -80.827
3.919 | -80.8265
3.9194 | -80.8261
3.9197 | -80.8252
3.92 | -80.8247
3.9204 | -80.8242
3.9207 | -80.8239
3.921 | -80.8237
3.9213 | -80.8236
3.9217 | -80.8233
3.922 | -80.8239
3.9223 | -80.8237
3.9227 | -80.8238
3.923 | -80.8235
3.9233 | -80.8232
3.9237 | -80.8238
3.924 | -80.8252
3.9243 | -80.826
3.9246 | -80.8263
3.925 | -80.8254
3.9253 | -80.8247
3.9256 | -80.8249
3.926 | -80.8247
3.9263 | -80.8245
3.9266 | -80.8245
3.9269 | -80.8243
3.9273 | -80.8231
3.9276 | -80.8201
3.9279 | -80.8187
3.9283 | -80.8188
3.9286 | -80.8182
3.9289 | -80.8176
3.9292 | -80.8174
3.9296 | -80.8173
3.9299 | -80.8173
3.9302 | -80.8165
3.9306 | -80.8165
3.9309 | -80.8179
3.9312 | -80.8189
3.9316 | -80.8206
3.9319 | -80.821
3.9322 | -80.8205
3.9325 | -80.8233
3.9329 | -80.8248
3.9332 | -80.8241
3.9335 | -80.8217
3.9339 | -80.8192
3.9342 | -80.8181
3.9345 | -80.8189
3.9348 | -80.8197
3.9352 | -80.8201
3.9355 | -80.8195
3.9358 | -80.819
3.9362 | -80.8187
3.9365 | -80.8196
3.9368 | -80.821
3.9371 | -80.821
3.9375 | -80.821
3.9378 | -80.8207
3.9381 | -80.8206
3.9385 | -80.8197
3.9388 | -80.8188
3.9391 | -80.8181
3.9395 | -80.8181
3.9398 | -80.8179
3.9401 | -80.8182
3.9404 | -80.8181
3.9408 | -80.8187
3.9411 | -80.8177
3.9414 | -80.817
3.9418 | -80.8171
3.9421 | -80.817
3.9424 | -80.8154
3.9427 | -80.8139
3.9431 | -80.8142
3.9434 | -80.8147
3.9437 | -80.8147
3.9441 | -80.8147
3.9444 | -80.8151
3.9447 | -80.8153
3.945 | -80.8151
3.9454 | -80.8147
3.9457 | -80.8137
3.946 | -80.8125
3.9464 | -80.8119
3.9467 | -80.8145
3.947 | -80.8144
3.9474 | -80.8141
3.9477 | -80.8145
3.948 | -80.8151
3.9483 | -80.8129
3.9487 | -80.8121
3.949 | -80.8121
3.9493 | -80.8116
3.9497 | -80.8116
3.95 | -80.8123
3.9503 | -80.8131
3.9506 | -80.8132
3.951 | -80.8131
3.9513 | -80.812
3.9516 | -80.8107
3.952 | -80.8095
3.9523 | -80.809
3.9526 | -80.8097
3.9529 | -80.8123
3.9533 | -80.8129
3.9536 | -80.812
3.9539 | -80.8106
3.9543 | -80.8096
3.9546 | -80.8086
3.9549 | -80.8094
3.9552 | -80.8099
3.9556 | -80.8105
3.9559 | -80.8103
3.9562 | -80.8102
3.9566 | -80.8103
3.9569 | -80.8101
3.9572 | -80.8092
3.9576 | -80.8079
3.9579 | -80.8082
3.9582 | -80.8086
3.9585 | -80.8087
3.9589 | -80.8082
3.9592 | -80.8066
3.9595 | -80.8063
3.9599 | -80.8074
3.9602 | -80.8082
3.9605 | -80.8098
3.9608 | -80.8115
3.9612 | -80.8117
3.9615 | -80.8117
3.9618 | -80.8118
3.9622 | -80.8118
3.9625 | -80.8122
3.9628 | -80.8134
3.9632 | -80.813
3.9635 | -80.813
3.9638 | -80.812
3.9641 | -80.8114
3.9645 | -80.8111
3.9648 | -80.8111
3.9651 | -80.8109
3.9655 | -80.8105
3.9658 | -80.8104
3.9661 | -80.8109
3.9664 | -80.8118
3.9668 | -80.8137
3.9671 | -80.8141
3.9674 | -80.815
3.9678 | -80.8155
3.9681 | -80.8153
3.9684 | -80.815
3.9687 | -80.8154
3.9691 | -80.8153
3.9694 | -80.8161
3.9697 | -80.8167
3.9701 | -80.8189
3.9704 | -80.8215
3.9707 | -80.8226
3.971 | -80.8228
3.9714 | -80.8229
3.9717 | -80.8229
3.972 | -80.8223
3.9724 | -80.8197
3.9727 | -80.8182
3.973 | -80.8179
3.9734 | -80.8195
3.9737 | -80.8204
3.974 | -80.8201
3.9743 | -80.8203
3.9747 | -80.8201
3.975 | -80.8199
3.9753 | -80.8213
3.9757 | -80.8228
3.976 | -80.8246
3.9763 | -80.8236
3.9766 | -80.8237
3.977 | -80.8237
3.9773 | -80.8237
3.9776 | -80.8253
3.978 | -80.8276
3.9783 | -80.8284
3.9786 | -80.8277
3.9789 | -80.826
3.9793 | -80.825
3.9796 | -80.8251
3.9799 | -80.826
3.9803 | -80.8257
3.9806 | -80.8253
3.9809 | -80.825
3.9813 | -80.8246
3.9816 | -80.8246
3.9819 | -80.8251
3.9822 | -80.8255
3.9826 | -80.8242
3.9829 | -80.8231
3.9832 | -80.823
3.9836 | -80.8228
3.9839 | -80.8224
3.9842 | -80.8219
3.9845 | -80.8208
3.9849 | -80.8201
3.9852 | -80.8208
3.9855 | -80.8211
3.9859 | -80.8197
3.9862 | -80.8183
3.9865 | -80.8179
3.9868 | -80.8179
3.9872 | -80.8182
3.9875 | -80.8177
3.9878 | -80.8181
3.9882 | -80.8188
3.9885 | -80.82
3.9888 | -80.82
3.9892 | -80.82
3.9895 | -80.8203
3.9898 | -80.8204
3.9901 | -80.8208
3.9905 | -80.8221
3.9908 | -80.824
3.9911 | -80.8244
3.9915 | -80.8229
3.9918 | -80.8229
3.9921 | -80.823
3.9924 | -80.8233
3.9928 | -80.8234
3.9931 | -80.824
3.9934 | -80.8244
3.9938 | -80.8252
3.9941 | -80.8256
3.9944 | -80.8268
3.9947 | -80.8288
3.9951 | -80.8296
3.9954 | -80.8301
3.9957 | -80.8291
3.9961 | -80.8283
3.9964 | -80.8295
3.9967 | -80.8285
3.9971 | -80.8265
3.9974 | -80.8259
3.9977 | -80.8263
3.998 | -80.8263
3.9984 | -80.8254
3.9987 | -80.8249
3.999 | -80.8244
3.9994 | -80.8239
3.9997 | -80.8226
4 | -80.8219
4.0003 | -80.8218
4.0007 | -80.8232
4.001 | -80.825
4.0013 | -80.8252
4.0017 | -80.8246
4.002 | -80.8248
4.0023 | -80.8254
4.0026 | -80.8256
4.003 | -80.826
4.0033 | -80.8276
4.0036 | -80.8289
4.004 | -80.8267
4.0043 | -80.8254
4.0046 | -80.8255
4.005 | -80.8254
4.0053 | -80.8256
4.0056 | -80.8258
4.0059 | -80.8252
4.0063 | -80.8241
4.0066 | -80.8242
4.0069 | -80.8245
4.0073 | -80.8241
4.0076 | -80.8249
4.0079 | -80.8253
4.0082 | -80.8256
4.0086 | -80.8258
4.0089 | -80.8257
4.0092 | -80.8256
4.0096 | -80.8258
4.0099 | -80.825
4.0102 | -80.8247
4.0105 | -80.8247
4.0109 | -80.8261
4.0112 | -80.829
4.0115 | -80.8306
4.0119 | -80.8311
4.0122 | -80.8315
4.0125 | -80.8308
4.0129 | -80.8304
4.0132 | -80.83
4.0135 | -80.8304
4.0138 | -80.8305
4.0142 | -80.8299
4.0145 | -80.8304
4.0148 | -80.8318
4.0152 | -80.8328
4.0155 | -80.8329
4.0158 | -80.8317
4.0161 | -80.831
4.0165 | -80.8313
4.0168 | -80.8322
4.0171 | -80.8326
4.0175 | -80.8317
4.0178 | -80.8313
4.0181 | -80.8315
4.0184 | -80.8316
4.0188 | -80.8316
4.0191 | -80.8313
4.0194 | -80.8303
4.0198 | -80.83
4.0201 | -80.8296
4.0204 | -80.8299
4.0207 | -80.83
4.0211 | -80.8311
4.0214 | -80.8326
4.0217 | -80.8332
4.0221 | -80.8345
4.0224 | -80.837
4.0227 | -80.8375
4.0231 | -80.8377
4.0234 | -80.8377
4.0237 | -80.8376
4.024 | -80.8375
4.0244 | -80.8373
4.0247 | -80.8373
4.025 | -80.8378
4.0254 | -80.838
4.0257 | -80.8382
4.026 | -80.8382
4.0263 | -80.8392
4.0267 | -80.8384
4.027 | -80.836
4.0273 | -80.8355
4.0277 | -80.8357
4.028 | -80.8359
4.0283 | -80.8361
4.0286 | -80.8362
4.029 | -80.836
4.0293 | -80.8352
4.0296 | -80.8343
4.03 | -80.8339
4.0303 | -80.8343
4.0306 | -80.8349
4.031 | -80.8346
4.0313 | -80.834
4.0316 | -80.834
4.0319 | -80.835
4.0323 | -80.8381
4.0326 | -80.8419
4.0329 | -80.843
4.0333 | -80.8434
4.0336 | -80.8439
4.0339 | -80.8439
4.0342 | -80.8432
4.0346 | -80.8444
4.0349 | -80.8451
4.0352 | -80.8453
4.0356 | -80.8467
4.0359 | -80.8472
4.0362 | -80.8466
4.0365 | -80.8476
4.0369 | -80.8477
4.0372 | -80.847
4.0375 | -80.8473
4.0379 | -80.848
4.0382 | -80.8479
4.0385 | -80.8479
4.0389 | -80.8483
4.0392 | -80.8487
4.0395 | -80.849
4.0398 | -80.8492
4.0402 | -80.8486
4.0405 | -80.8488
4.0408 | -80.8501
4.0412 | -80.8507
4.0415 | -80.8502
4.0418 | -80.8496
4.0421 | -80.8489
4.0425 | -80.8477
4.0428 | -80.8448
4.0431 | -80.8437
4.0435 | -80.844
4.0438 | -80.8441
4.0441 | -80.8438
4.0444 | -80.8438
4.0448 | -80.8437
4.0451 | -80.8458
4.0454 | -80.8464
4.0458 | -80.8484
4.0461 | -80.8486
4.0464 | -80.8478
4.0468 | -80.8481
4.0471 | -80.8488
4.0474 | -80.8481
4.0477 | -80.8476
4.0481 | -80.8475
4.0484 | -80.8477
4.0487 | -80.8478
4.0491 | -80.8486
4.0494 | -80.8497
4.0497 | -80.8504
4.05 | -80.8506
4.0504 | -80.8507
4.0507 | -80.8505
4.051 | -80.8501
4.0514 | -80.8485
4.0517 | -80.8474
4.052 | -80.8456
4.0523 | -80.8431
4.0527 | -80.8421
4.053 | -80.8417
4.0533 | -80.842
4.0537 | -80.8419
4.054 | -80.8404
4.0543 | -80.8392
4.0546 | -80.8382
4.055 | -80.8383
4.0553 | -80.8391
4.0556 | -80.8399
4.056 | -80.8405
4.0563 | -80.8407
4.0566 | -80.8402
4.057 | -80.8394
4.0573 | -80.8394
4.0576 | -80.8392
4.0579 | -80.8389
4.0583 | -80.8389
4.0586 | -80.8387
4.0589 | -80.8392
4.0593 | -80.8391
4.0596 | -80.8392
4.0599 | -80.8393
4.0602 | -80.8387
4.0606 | -80.8386
4.0609 | -80.8379
4.0612 | -80.8387
4.0616 | -80.8395
4.0619 | -80.8398
4.0622 | -80.8413
4.0625 | -80.8432
4.0629 | -80.8432
4.0632 | -80.8433
4.0635 | -80.8448
4.0639 | -80.8483
4.0642 | -80.8483
4.0645 | -80.8478
4.0649 | -80.8476
4.0652 | -80.8472
4.0655 | -80.8469
4.0658 | -80.8462
4.0662 | -80.846
4.0665 | -80.8455
4.0668 | -80.8449
4.0672 | -80.8446
4.0675 | -80.8464
4.0678 | -80.8471
4.0681 | -80.8475
4.0685 | -80.8501
4.0688 | -80.8518
4.0691 | -80.851
4.0695 | -80.8505
4.0698 | -80.8497
4.0701 | -80.8486
4.0704 | -80.8485
4.0708 | -80.8485
4.0711 | -80.8482
4.0714 | -80.8481
4.0718 | -80.8463
4.0721 | -80.8442
4.0724 | -80.8441
4.0728 | -80.8436
4.0731 | -80.8432
4.0734 | -80.8433
4.0737 | -80.8435
4.0741 | -80.8445
4.0744 | -80.846
4.0747 | -80.8456
4.0751 | -80.8452
4.0754 | -80.8444
4.0757 | -80.844
4.076 | -80.844
4.0764 | -80.8438
4.0767 | -80.8443
4.077 | -80.8448
4.0774 | -80.8446
4.0777 | -80.8445
4.078 | -80.8454
4.0783 | -80.8465
4.0787 | -80.8468
4.079 | -80.8475
4.0793 | -80.8483
4.0797 | -80.8486
4.08 | -80.849
4.0803 | -80.8499
4.0807 | -80.8498
4.081 | -80.8498
4.0813 | -80.8497
4.0816 | -80.8504
4.082 | -80.8515
4.0823 | -80.8529
4.0826 | -80.8533
4.083 | -80.8514
4.0833 | -80.8519
4.0836 | -80.8518
4.0839 | -80.8517
4.0843 | -80.8501
4.0846 | -80.8489
4.0849 | -80.8488
4.0853 | -80.8496
4.0856 | -80.8506
4.0859 | -80.8511
4.0862 | -80.8503
4.0866 | -80.85
4.0869 | -80.8496
4.0872 | -80.8492
4.0876 | -80.8494
4.0879 | -80.8493
4.0882 | -80.8491
4.0886 | -80.8494
4.0889 | -80.8497
4.0892 | -80.8496
4.0895 | -80.8496
4.0899 | -80.8495
4.0902 | -80.8491
4.0905 | -80.8485
4.0909 | -80.8495
4.0912 | -80.8515
4.0915 | -80.8525
4.0918 | -80.8529
4.0922 | -80.8518
4.0925 | -80.8497
4.0928 | -80.8488
4.0932 | -80.8473
4.0935 | -80.8482
4.0938 | -80.8492
4.0941 | -80.8489
4.0945 | -80.8485
4.0948 | -80.8482
4.0951 | -80.8485
4.0955 | -80.8487
4.0958 | -80.8485
4.0961 | -80.8486
4.0965 | -80.8481
4.0968 | -80.846
4.0971 | -80.8461
4.0974 | -80.8472
4.0978 | -80.8461
4.0981 | -80.8436
4.0984 | -80.8418
4.0988 | -80.8417
4.0991 | -80.8414
4.0994 | -80.8412
4.0997 | -80.8418
4.1001 | -80.8427
4.1004 | -80.8417
4.1007 | -80.8387
4.1011 | -80.8376
4.1014 | -80.8374
4.1017 | -80.8372
4.102 | -80.837
4.1024 | -80.8372
4.1027 | -80.8374
4.103 | -80.8369
4.1034 | -80.8372
4.1037 | -80.8369
4.104 | -80.8374
4.1044 | -80.8372
4.1047 | -80.8375
4.105 | -80.837
4.1053 | -80.8373
4.1057 | -80.8376
4.106 | -80.8374
4.1063 | -80.8369
4.1067 | -80.8367
4.107 | -80.8369
4.1073 | -80.837
4.1076 | -80.8387
4.108 | -80.8389
4.1083 | -80.8378
4.1086 | -80.8377
4.109 | -80.838
4.1093 | -80.8397
4.1096 | -80.8421
4.1099 | -80.8423
4.1103 | -80.8419
4.1106 | -80.8419
4.1109 | -80.8418
4.1113 | -80.8424
4.1116 | -80.842
4.1119 | -80.8421
4.1122 | -80.8426
4.1126 | -80.8429
4.1129 | -80.8422
4.1132 | -80.842
4.1136 | -80.8421
4.1139 | -80.8424
4.1142 | -80.8415
4.1146 | -80.8402
4.1149 | -80.8397
4.1152 | -80.8403
4.1155 | -80.841
4.1159 | -80.84
4.1162 | -80.8376
4.1165 | -80.8366
4.1169 | -80.8369
4.1172 | -80.8373
4.1175 | -80.8374
4.1178 | -80.8375
4.1182 | -80.8375
4.1185 | -80.8366
4.1188 | -80.836
4.1192 | -80.8353
4.1195 | -80.8353
4.1198 | -80.8354
4.1201 | -80.8351
4.1205 | -80.8355
4.1208 | -80.8359
4.1211 | -80.8371
4.1215 | -80.8365
4.1218 | -80.8354
4.1221 | -80.8352
4.1225 | -80.8333
4.1228 | -80.8362
4.1231 | -80.8379
4.1234 | -80.838
4.1238 | -80.838
4.1241 | -80.8373
4.1244 | -80.8375
4.1248 | -80.8376
4.1251 | -80.8374
4.1254 | -80.8371
4.1257 | -80.8362
4.1261 | -80.8358
4.1264 | -80.8368
4.1267 | -80.8377
4.1271 | -80.8376
4.1274 | -80.8373
4.1277 | -80.8379
4.128 | -80.8389
4.1284 | -80.8389
4.1287 | -80.8391
4.129 | -80.8391
4.1294 | -80.8394
4.1297 | -80.8397
4.13 | -80.84
4.1304 | -80.8395
4.1307 | -80.8406
4.131 | -80.8408
4.1313 | -80.8413
4.1317 | -80.8412
4.132 | -80.8414
4.1323 | -80.8423
4.1327 | -80.8427
4.133 | -80.8422
4.1333 | -80.8432
4.1336 | -80.8447
4.134 | -80.8451
4.1343 | -80.845
4.1346 | -80.846
4.135 | -80.8469
4.1353 | -80.8471
4.1356 | -80.8471
4.1359 | -80.8471
4.1363 | -80.8475
4.1366 | -80.8477
4.1369 | -80.8485
4.1373 | -80.8491
4.1376 | -80.8491
4.1379 | -80.8484
4.1383 | -80.8489
4.1386 | -80.8488
4.1389 | -80.8491
4.1392 | -80.8505
4.1396 | -80.8516
4.1399 | -80.8517
4.1402 | -80.8527
4.1406 | -80.8528
4.1409 | -80.8533
4.1412 | -80.853
4.1415 | -80.8534
4.1419 | -80.8534
4.1422 | -80.8544
4.1425 | -80.8562
4.1429 | -80.8572
4.1432 | -80.8588
4.1435 | -80.8614
4.1438 | -80.8635
4.1442 | -80.8644
4.1445 | -80.865
4.1448 | -80.8637
4.1452 | -80.8644
4.1455 | -80.8667
4.1458 | -80.868
4.1461 | -80.8695
4.1465 | -80.8703
4.1468 | -80.8711
4.1471 | -80.8729
4.1475 | -80.8749
4.1478 | -80.8792
4.1481 | -80.8801
4.1485 | -80.8801
4.1488 | -80.8814
4.1491 | -80.8814
4.1494 | -80.8816
4.1498 | -80.8829
4.1501 | -80.8843
4.1504 | -80.885
4.1508 | -80.8853
4.1511 | -80.885
4.1514 | -80.8843
4.1517 | -80.8849
4.1521 | -80.8858
4.1524 | -80.8869
4.1527 | -80.8869
4.1531 | -80.8868
4.1534 | -80.8872
4.1537 | -80.8883
4.154 | -80.8886
4.1544 | -80.8888
4.1547 | -80.889
4.155 | -80.8902
4.1554 | -80.8909
4.1557 | -80.8909
4.156 | -80.8907
4.1564 | -80.8904
4.1567 | -80.8912
4.157 | -80.8921
4.1573 | -80.892
4.1577 | -80.8908
4.158 | -80.8907
4.1583 | -80.8913
4.1587 | -80.8917
4.159 | -80.893
4.1593 | -80.8946
4.1596 | -80.8959
4.16 | -80.8959
4.1603 | -80.8965
4.1606 | -80.8967
4.161 | -80.8961
4.1613 | -80.8964
4.1616 | -80.897
4.1619 | -80.8972
4.1623 | -80.8971
4.1626 | -80.8961
4.1629 | -80.8962
4.1633 | -80.8956
4.1636 | -80.8962
4.1639 | -80.8967
4.1643 | -80.8977
4.1646 | -80.8983
4.1649 | -80.899
4.1652 | -80.8987
4.1656 | -80.8984
4.1659 | -80.8995
4.1662 | -80.9013
4.1666 | -80.902
4.1669 | -80.9021
4.1672 | -80.9016
4.1675 | -80.9015
4.1679 | -80.9024
4.1682 | -80.9032
4.1685 | -80.9046
4.1689 | -80.9061
4.1692 | -80.9077
4.1695 | -80.9084
4.1698 | -80.9091
4.1702 | -80.9103
4.1705 | -80.9107
4.1708 | -80.91
4.1712 | -80.9102
4.1715 | -80.9103
4.1718 | -80.9111
4.1722 | -80.9112
4.1725 | -80.9117
4.1728 | -80.9123
4.1731 | -80.913
4.1735 | -80.9144
4.1738 | -80.9149
4.1741 | -80.9153
4.1745 | -80.9158
4.1748 | -80.9162
4.1751 | -80.9165
4.1754 | -80.9174
4.1758 | -80.9181
4.1761 | -80.9183
4.1764 | -80.9191
4.1768 | -80.9208
4.1771 | -80.9218
4.1774 | -80.9226
4.1777 | -80.9235
4.1781 | -80.924
4.1784 | -80.9237
4.1787 | -80.9243
4.1791 | -80.9253
4.1794 | -80.9257
4.1797 | -80.9261
4.1801 | -80.9271
4.1804 | -80.9277
4.1807 | -80.9271
4.181 | -80.9259
4.1814 | -80.9261
4.1817 | -80.9247
4.182 | -80.9237
4.1824 | -80.9247
4.1827 | -80.9256
4.183 | -80.9263
4.1833 | -80.9256
4.1837 | -80.9259
4.184 | -80.9267
4.1843 | -80.9279
4.1847 | -80.9295
4.185 | -80.9302
4.1853 | -80.9309
4.1856 | -80.9314
4.186 | -80.9319
4.1863 | -80.9325
4.1866 | -80.933
4.187 | -80.933
4.1873 | -80.9336
4.1876 | -80.9342
4.188 | -80.9348
4.1883 | -80.9354
4.1886 | -80.9346
4.1889 | -80.9345
4.1893 | -80.9349
4.1896 | -80.9353
4.1899 | -80.9358
4.1903 | -80.9362
4.1906 | -80.9367
4.1909 | -80.9376
4.1912 | -80.9384
4.1916 | -80.9394
4.1919 | -80.9417
4.1922 | -80.944
4.1926 | -80.9447
4.1929 | -80.9455
4.1932 | -80.9458
4.1935 | -80.9465
4.1939 | -80.9467
4.1942 | -80.9472
4.1945 | -80.9473
4.1949 | -80.9479
4.1952 | -80.9495
4.1955 | -80.9499
4.1959 | -80.9491
4.1962 | -80.9493
4.1965 | -80.9499
4.1968 | -80.9498
4.1972 | -80.9488
4.1975 | -80.9493
4.1978 | -80.9511
4.1982 | -80.9507
4.1985 | -80.9515
4.1988 | -80.9529
4.1991 | -80.9534
4.1995 | -80.9541
4.1998 | -80.9547
4.2001 | -80.956
4.2005 | -80.9578
4.2008 | -80.9584
4.2011 | -80.9584
4.2014 | -80.9572
4.2018 | -80.9557
4.2021 | -80.9557
4.2024 | -80.9563
4.2028 | -80.9562
4.2031 | -80.9566
4.2034 | -80.9568
4.2038 | -80.9575
4.2041 | -80.9583
4.2044 | -80.9591
4.2047 | -80.9599
4.2051 | -80.9608
4.2054 | -80.9613
4.2057 | -80.9622
4.2061 | -80.9629
4.2064 | -80.9634
4.2067 | -80.9636
4.207 | -80.9622
4.2074 | -80.9621
4.2077 | -80.9633
4.208 | -80.9638
4.2084 | -80.9645
4.2087 | -80.9655
4.209 | -80.9661
4.2093 | -80.9662
4.2097 | -80.9668
4.21 | -80.9684
4.2103 | -80.9713
4.2107 | -80.973
4.211 | -80.9732
4.2113 | -80.974
4.2116 | -80.975
4.212 | -80.9756
4.2123 | -80.976
4.2126 | -80.9754
4.213 | -80.9764
4.2133 | -80.9773
4.2136 | -80.9779
4.214 | -80.9784
4.2143 | -80.9786
4.2146 | -80.9791
4.2149 | -80.9794
4.2153 | -80.9811
4.2156 | -80.9813
4.2159 | -80.982
4.2163 | -80.983
4.2166 | -80.983
4.2169 | -80.9857
4.2172 | -80.9903
4.2176 | -80.9916
4.2179 | -80.992
4.2182 | -80.9922
4.2186 | -80.9927
4.2189 | -80.9935
4.2192 | -80.993
4.2195 | -80.9921
4.2199 | -80.9922
4.2202 | -80.9927
4.2205 | -80.9931
4.2209 | -80.9938
4.2212 | -80.9947
4.2215 | -80.9955
4.2219 | -80.9967
4.2222 | -80.9978
4.2225 | -80.9988
4.2228 | -80.9991
4.2232 | -80.9991
4.2235 | -80.9996
4.2238 | -80.9995
4.2242 | -80.9989
4.2245 | -80.9988
4.2248 | -81.0003
4.2251 | -81.0015
4.2255 | -81.0026
4.2258 | -81.0038
4.2261 | -81.0046
4.2265 | -81.0062
4.2268 | -81.007
4.2271 | -81.0074
4.2274 | -81.0079
4.2278 | -81.008
4.2281 | -81.0085
4.2284 | -81.0088
4.2288 | -81.0089
4.2291 | -81.0094
4.2294 | -81.0107
4.2298 | -81.0126
4.2301 | -81.0148
4.2304 | -81.0154
4.2307 | -81.014
4.2311 | -81.0127
4.2314 | -81.0141
4.2317 | -81.0154
4.2321 | -81.0156
4.2324 | -81.0175
4.2327 | -81.0188
4.233 | -81.0191
4.2334 | -81.0199
4.2337 | -81.0194
4.234 | -81.0183
4.2344 | -81.0183
4.2347 | -81.0187
4.235 | -81.0194
4.2353 | -81.0196
4.2357 | -81.0197
4.236 | -81.0206
4.2363 | -81.0223
4.2367 | -81.0238
4.237 | -81.0242
4.2373 | -81.0246
4.2377 | -81.0248
4.238 | -81.0254
4.2383 | -81.0268
4.2386 | -81.0282
4.239 | -81.0272
4.2393 | -81.0271
4.2396 | -81.0268
4.24 | -81.0264
4.2403 | -81.0269
4.2406 | -81.027
4.2409 | -81.027
4.2413 | -81.0274
4.2416 | -81.028
4.2419 | -81.0286
4.2423 | -81.0306
4.2426 | -81.0311
4.2429 | -81.0305
4.2432 | -81.03
4.2436 | -81.0299
4.2439 | -81.0311
4.2442 | -81.0315
4.2446 | -81.0317
4.2449 | -81.032
4.2452 | -81.0318
4.2455 | -81.0315
4.2459 | -81.0309
4.2462 | -81.0295
4.2465 | -81.0276
4.2469 | -81.0266
4.2472 | -81.0265
4.2475 | -81.0267
4.2479 | -81.0271
4.2482 | -81.0274
4.2485 | -81.0283
4.2488 | -81.0286
4.2492 | -81.0298
4.2495 | -81.0305
4.2498 | -81.0298
4.2502 | -81.0294
4.2505 | -81.0296
4.2508 | -81.0313
4.2511 | -81.0317
4.2515 | -81.0316
4.2518 | -81.0316
4.2521 | -81.0318
4.2525 | -81.0306
4.2528 | -81.0306
4.2531 | -81.0311
4.2534 | -81.0317
4.2538 | -81.0327
4.2541 | -81.0325
4.2544 | -81.0325
4.2548 | -81.0329
4.2551 | -81.0335
4.2554 | -81.0332
4.2558 | -81.0335
4.2561 | -81.0334
4.2564 | -81.0326
4.2567 | -81.032
4.2571 | -81.0317
4.2574 | -81.0318
4.2577 | -81.0316
4.2581 | -81.0325
4.2584 | -81.0352
4.2587 | -81.037
4.259 | -81.0379
4.2594 | -81.0378
4.2597 | -81.0372
4.26 | -81.0376
4.2604 | -81.0372
4.2607 | -81.037
4.261 | -81.0381
4.2613 | -81.0375
4.2617 | -81.0372
4.262 | -81.0375
4.2623 | -81.0389
4.2627 | -81.039
4.263 | -81.0392
4.2633 | -81.0394
4.2637 | -81.0355
4.264 | -81.0353
4.2643 | -81.0354
4.2646 | -81.0349
4.265 | -81.0357
4.2653 | -81.0371
4.2656 | -81.0388
4.266 | -81.0403
4.2663 | -81.0405
4.2666 | -81.0403
4.2669 | -81.0408
4.2673 | -81.0407
4.2676 | -81.041
4.2679 | -81.0413
4.2683 | -81.0408
4.2686 | -81.0404
4.2689 | -81.0404
4.2692 | -81.0418
4.2696 | -81.0429
4.2699 | -81.0412
4.2702 | -81.0393
4.2706 | -81.0394
4.2709 | -81.0392
4.2712 | -81.039
4.2716 | -81.0396
4.2719 | -81.0396
4.2722 | -81.0395
4.2725 | -81.041
4.2729 | -81.0425
4.2732 | -81.0433
4.2735 | -81.044
4.2739 | -81.0437
4.2742 | -81.0415
4.2745 | -81.0418
4.2748 | -81.0422
4.2752 | -81.0422
4.2755 | -81.0417
4.2758 | -81.0415
4.2762 | -81.0411
4.2765 | -81.0407
4.2768 | -81.0403
4.2771 | -81.0403
4.2775 | -81.0401
4.2778 | -81.0398
4.2781 | -81.0396
4.2785 | -81.0406
4.2788 | -81.043
4.2791 | -81.0433
4.2795 | -81.0431
4.2798 | -81.043
4.2801 | -81.0423
4.2804 | -81.0409
4.2808 | -81.0401
4.2811 | -81.0402
4.2814 | -81.0391
4.2818 | -81.0382
4.2821 | -81.037
4.2824 | -81.0365
4.2827 | -81.0363
4.2831 | -81.0357
4.2834 | -81.0365
4.2837 | -81.0374
4.2841 | -81.0374
4.2844 | -81.0387
4.2847 | -81.0406
4.285 | -81.0411
4.2854 | -81.0402
4.2857 | -81.0394
4.286 | -81.0393
4.2864 | -81.0401
4.2867 | -81.0411
4.287 | -81.0415
4.2874 | -81.0421
4.2877 | -81.0435
4.288 | -81.044
4.2883 | -81.0444
4.2887 | -81.0445
4.289 | -81.0441
4.2893 | -81.0435
4.2897 | -81.044
4.29 | -81.0438
4.2903 | -81.0436
4.2906 | -81.0435
4.291 | -81.0436
4.2913 | -81.0438
4.2916 | -81.0438
4.292 | -81.0448
4.2923 | -81.044
4.2926 | -81.0453
4.2929 | -81.0446
4.2933 | -81.0442
4.2936 | -81.0439
4.2939 | -81.0438
4.2943 | -81.045
4.2946 | -81.0454
4.2949 | -81.0439
4.2953 | -81.0424
4.2956 | -81.0422
4.2959 | -81.0427
4.2962 | -81.0422
4.2966 | -81.0425
4.2969 | -81.0423
4.2972 | -81.0416
4.2976 | -81.0414
4.2979 | -81.042
4.2982 | -81.0415
4.2985 | -81.0423
4.2989 | -81.0426
4.2992 | -81.0428
4.2995 | -81.0418
4.2999 | -81.0414
4.3002 | -81.0419
4.3005 | -81.0421
4.3008 | -81.0417
4.3012 | -81.0426
4.3015 | -81.0429
4.3018 | -81.042
4.3022 | -81.0415
4.3025 | -81.0415
4.3028 | -81.0414
4.3031 | -81.0417
4.3035 | -81.0413
4.3038 | -81.0409
4.3041 | -81.0405
4.3045 | -81.04
4.3048 | -81.0409
4.3051 | -81.0419
4.3055 | -81.0419
4.3058 | -81.0417
4.3061 | -81.0413
4.3064 | -81.0415
4.3068 | -81.0428
4.3071 | -81.0432
4.3074 | -81.0432
4.3078 | -81.0428
4.3081 | -81.0427
4.3084 | -81.0427
4.3087 | -81.0431
4.3091 | -81.0443
4.3094 | -81.0442
4.3097 | -81.0439
4.3101 | -81.0424
4.3104 | -81.0405
4.3107 | -81.0389
4.311 | -81.0375
4.3114 | -81.0374
4.3117 | -81.036
4.312 | -81.0356
4.3124 | -81.0365
4.3127 | -81.0362
4.313 | -81.0358
4.3134 | -81.036
4.3137 | -81.0361
4.314 | -81.0366
4.3143 | -81.0365
4.3147 | -81.0362
4.315 | -81.0362
4.3153 | -81.036
4.3157 | -81.0356
4.316 | -81.0356
4.3163 | -81.0337
4.3166 | -81.0336
4.317 | -81.036
4.3173 | -81.0362
4.3176 | -81.0352
4.318 | -81.0345
4.3183 | -81.0344
4.3186 | -81.0331
4.3189 | -81.0307
4.3193 | -81.0317
4.3196 | -81.0334
4.3199 | -81.0335
4.3203 | -81.0331
4.3206 | -81.0329
4.3209 | -81.0326
4.3213 | -81.0327
4.3216 | -81.0332
4.3219 | -81.0337
4.3222 | -81.0337
4.3226 | -81.0334
4.3229 | -81.0328
4.3232 | -81.0325
4.3236 | -81.0314
4.3239 | -81.0297
4.3242 | -81.0283
4.3245 | -81.0293
4.3249 | -81.0294
4.3252 | -81.0292
4.3255 | -81.0285
4.3259 | -81.0278
4.3262 | -81.0278
4.3265 | -81.0282
4.3268 | -81.0287
4.3272 | -81.0279
4.3275 | -81.0272
4.3278 | -81.0274
4.3282 | -81.0272
4.3285 | -81.0267
4.3288 | -81.0268
4.3292 | -81.0266
4.3295 | -81.026
4.3298 | -81.0256
4.3301 | -81.0261
4.3305 | -81.0265
4.3308 | -81.0264
4.3311 | -81.027
4.3315 | -81.027
4.3318 | -81.0261
4.3321 | -81.0246
4.3324 | -81.0254
4.3328 | -81.026
4.3331 | -81.0269
4.3334 | -81.0284
4.3338 | -81.0311
4.3341 | -81.0318
4.3344 | -81.0313
4.3347 | -81.0313
4.3351 | -81.0317
4.3354 | -81.0311
4.3357 | -81.0313
4.3361 | -81.0315
4.3364 | -81.0317
4.3367 | -81.0315
4.337 | -81.0314
4.3374 | -81.0313
4.3377 | -81.032
4.338 | -81.0319
4.3384 | -81.03
4.3387 | -81.0312
4.339 | -81.032
4.3394 | -81.0319
4.3397 | -81.0317
4.34 | -81.032
4.3403 | -81.032
4.3407 | -81.0322
4.341 | -81.0323
4.3413 | -81.0312
4.3417 | -81.0311
4.342 | -81.0313
4.3423 | -81.0317
4.3426 | -81.032
4.343 | -81.0338
4.3433 | -81.0352
4.3436 | -81.0372
4.344 | -81.0393
4.3443 | -81.0404
4.3446 | -81.0402
4.3449 | -81.0402
4.3453 | -81.0404
4.3456 | -81.0416
4.3459 | -81.0427
4.3463 | -81.0436
4.3466 | -81.0442
4.3469 | -81.0441
4.3473 | -81.044
4.3476 | -81.0437
4.3479 | -81.0438
4.3482 | -81.0441
4.3486 | -81.0444
4.3489 | -81.0451
4.3492 | -81.0463
4.3496 | -81.0451
4.3499 | -81.0452
4.3502 | -81.0454
4.3505 | -81.0447
4.3509 | -81.0447
4.3512 | -81.043
4.3515 | -81.0417
4.3519 | -81.0415
4.3522 | -81.0416
4.3525 | -81.0421
4.3528 | -81.0427
4.3532 | -81.0433
4.3535 | -81.0431
4.3538 | -81.0434
4.3542 | -81.0433
4.3545 | -81.0434
4.3548 | -81.044
4.3552 | -81.0444
4.3555 | -81.0444
4.3558 | -81.0447
4.3561 | -81.0448
4.3565 | -81.0447
4.3568 | -81.0447
4.3571 | -81.0454
4.3575 | -81.0459
4.3578 | -81.0463
4.3581 | -81.0455
4.3584 | -81.046
4.3588 | -81.0471
4.3591 | -81.0473
4.3594 | -81.0478
4.3598 | -81.0481
4.3601 | -81.0474
4.3604 | -81.0474
4.3607 | -81.0473
4.3611 | -81.0471
4.3614 | -81.0476
4.3617 | -81.0492
4.3621 | -81.0504
4.3624 | -81.051
4.3627 | -81.0519
4.3631 | -81.053
4.3634 | -81.0532
4.3637 | -81.0536
4.364 | -81.0539
4.3644 | -81.0538
4.3647 | -81.0555
4.365 | -81.0575
4.3654 | -81.0583
4.3657 | -81.0591
4.366 | -81.0597
4.3663 | -81.0597
4.3667 | -81.0595
4.367 | -81.0591
4.3673 | -81.0595
4.3677 | -81.0599
4.368 | -81.0602
4.3683 | -81.0611
4.3686 | -81.0618
4.369 | -81.0626
4.3693 | -81.0624
4.3696 | -81.0625
4.37 | -81.0634
4.3703 | -81.0638
4.3706 | -81.0639
4.371 | -81.0642
4.3713 | -81.0642
4.3716 | -81.0639
4.3719 | -81.0638
4.3723 | -81.0638
4.3726 | -81.0642
4.3729 | -81.0646
4.3733 | -81.0646
4.3736 | -81.0642
4.3739 | -81.0643
4.3742 | -81.0649
4.3746 | -81.0653
4.3749 | -81.0658
4.3752 | -81.0665
4.3756 | -81.0671
4.3759 | -81.0674
4.3762 | -81.0674
4.3765 | -81.0695
4.3769 | -81.0697
4.3772 | -81.0705
4.3775 | -81.0713
4.3779 | -81.0707
4.3782 | -81.0711
4.3785 | -81.0736
4.3789 | -81.0755
4.3792 | -81.0764
4.3795 | -81.0761
4.3798 | -81.0765
4.3802 | -81.077
4.3805 | -81.0784
4.3808 | -81.0798
4.3812 | -81.0808
4.3815 | -81.0809
4.3818 | -81.0813
4.3821 | -81.0818
4.3825 | -81.0827
4.3828 | -81.0822
4.3831 | -81.0821
4.3835 | -81.0828
4.3838 | -81.0832
4.3841 | -81.0835
4.3844 | -81.0838
4.3848 | -81.0846
4.3851 | -81.0848
4.3854 | -81.085
4.3858 | -81.0856
4.3861 | -81.0862
4.3864 | -81.0881
4.3868 | -81.09
4.3871 | -81.0911
4.3874 | -81.0923
4.3877 | -81.0933
4.3881 | -81.0957
4.3884 | -81.0961
4.3887 | -81.0969
4.3891 | -81.0979
4.3894 | -81.0988
4.3897 | -81.0996
4.39 | -81.1
4.3904 | -81.1005
4.3907 | -81.1006
4.391 | -81.1009
4.3914 | -81.1024
4.3917 | -81.1041
4.392 | -81.1051
4.3923 | -81.106
4.3927 | -81.1075
4.393 | -81.1101
4.3933 | -81.1122
4.3937 | -81.1124
4.394 | -81.1127
4.3943 | -81.1131
4.3947 | -81.114
4.395 | -81.1147
4.3953 | -81.1165
4.3956 | -81.1202
4.396 | -81.1221
4.3963 | -81.1227
4.3966 | -81.1247
4.397 | -81.1257
4.3973 | -81.1261
4.3976 | -81.1268
4.3979 | -81.1275
4.3983 | -81.1282
4.3986 | -81.1274
4.3989 | -81.1259
4.3993 | -81.1255
4.3996 | -81.1257
4.3999 | -81.1264
4.4002 | -81.1262
4.4006 | -81.1254
4.4009 | -81.1254
4.4012 | -81.1261
4.4016 | -81.1269
4.4019 | -81.1281
4.4022 | -81.1291
4.4025 | -81.1301
4.4029 | -81.1304
4.4032 | -81.1303
4.4035 | -81.1313
4.4039 | -81.1328
4.4042 | -81.1341
4.4045 | -81.1349
4.4049 | -81.1362
4.4052 | -81.1371
4.4055 | -81.1384
4.4058 | -81.1396
4.4062 | -81.141
4.4065 | -81.1417
4.4068 | -81.1423
4.4072 | -81.1431
4.4075 | -81.1445
4.4078 | -81.1448
4.4081 | -81.1449
4.4085 | -81.1453
4.4088 | -81.1462
4.4091 | -81.1476
4.4095 | -81.1485
4.4098 | -81.1498
4.4101 | -81.1506
4.4104 | -81.1528
4.4108 | -81.1548
4.4111 | -81.1566
4.4114 | -81.1576
4.4118 | -81.1585
4.4121 | -81.1612
4.4124 | -81.1639
4.4128 | -81.1643
4.4131 | -81.1638
4.4134 | -81.1642
4.4137 | -81.1651
4.4141 | -81.165
4.4144 | -81.1652
4.4147 | -81.1665
4.4151 | -81.1686
4.4154 | -81.1702
4.4157 | -81.17
4.416 | -81.1706
4.4164 | -81.1705
4.4167 | -81.1709
4.417 | -81.1711
4.4174 | -81.172
4.4177 | -81.1724
4.418 | -81.1729
4.4183 | -81.1744
4.4187 | -81.1764
4.419 | -81.1774
4.4193 | -81.1812
4.4197 | -81.1833
4.42 | -81.1837
4.4203 | -81.1846
4.4207 | -81.1855
4.421 | -81.1871
4.4213 | -81.1886
4.4216 | -81.189
4.422 | -81.1886
4.4223 | -81.1895
4.4226 | -81.1911
4.423 | -81.1911
4.4233 | -81.1915
4.4236 | -81.1923
4.4239 | -81.1933
4.4243 | -81.1946
4.4246 | -81.1952
4.4249 | -81.1965
4.4253 | -81.1984
4.4256 | -81.1994
4.4259 | -81.1998
4.4262 | -81.2009
4.4266 | -81.2015
4.4269 | -81.2024
4.4272 | -81.2035
4.4276 | -81.2048
4.4279 | -81.2046
4.4282 | -81.2052
4.4286 | -81.2051
4.4289 | -81.2057
4.4292 | -81.2069
4.4295 | -81.2089
4.4299 | -81.2099
4.4302 | -81.2104
4.4305 | -81.2117
4.4309 | -81.213
4.4312 | -81.2142
4.4315 | -81.215
4.4318 | -81.2162
4.4322 | -81.2174
4.4325 | -81.2182
4.4328 | -81.2193
4.4332 | -81.2206
4.4335 | -81.2219
4.4338 | -81.2228
4.4341 | -81.223
4.4345 | -81.2221
4.4348 | -81.223
4.4351 | -81.224
4.4355 | -81.2253
4.4358 | -81.2268
4.4361 | -81.2278
4.4364 | -81.229
4.4368 | -81.2299
4.4371 | -81.2308
4.4374 | -81.2319
4.4378 | -81.2328
4.4381 | -81.2337
4.4384 | -81.2352
4.4388 | -81.2362
4.4391 | -81.2371
4.4394 | -81.2392
4.4397 | -81.2404
4.4401 | -81.2417
4.4404 | -81.2425
4.4407 | -81.2431
4.4411 | -81.2441
4.4414 | -81.245
4.4417 | -81.2461
4.442 | -81.2469
4.4424 | -81.2478
4.4427 | -81.2492
4.443 | -81.2516
4.4434 | -81.2512
4.4437 | -81.2488
4.444 | -81.2486
4.4443 | -81.2486
4.4447 | -81.2495
4.445 | -81.2496
4.4453 | -81.2496
4.4457 | -81.2504
4.446 | -81.2508
4.4463 | -81.2522
4.4467 | -81.2535
4.447 | -81.2548
4.4473 | -81.2558
4.4476 | -81.2575
4.448 | -81.2577
4.4483 | -81.2584
4.4486 | -81.2592
4.449 | -81.2583
4.4493 | -81.2592
4.4496 | -81.2601
4.4499 | -81.2615
4.4503 | -81.2648
4.4506 | -81.2667
4.4509 | -81.2675
4.4513 | -81.2677
4.4516 | -81.2679
4.4519 | -81.2693
4.4522 | -81.2703
4.4526 | -81.2709
4.4529 | -81.2721
4.4532 | -81.2728
4.4536 | -81.2743
4.4539 | -81.2752
4.4542 | -81.2765
4.4546 | -81.2776
4.4549 | -81.2787
4.4552 | -81.2801
4.4555 | -81.2819
4.4559 | -81.2828
4.4562 | -81.2821
4.4565 | -81.2846
4.4569 | -81.2877
4.4572 | -81.2888
4.4575 | -81.2899
4.4578 | -81.2911
4.4582 | -81.292
4.4585 | -81.2931
4.4588 | -81.2933
4.4592 | -81.2939
4.4595 | -81.2953
4.4598 | -81.2969
4.4601 | -81.2982
4.4605 | -81.3002
4.4608 | -81.3012
4.4611 | -81.3023
4.4615 | -81.3034
4.4618 | -81.3053
4.4621 | -81.3073
4.4625 | -81.3079
4.4628 | -81.31
4.4631 | -81.3117
4.4634 | -81.3128
4.4638 | -81.3129
4.4641 | -81.3128
4.4644 | -81.3137
4.4648 | -81.3146
4.4651 | -81.3161
4.4654 | -81.3156
4.4657 | -81.3162
4.4661 | -81.3174
4.4664 | -81.3195
4.4667 | -81.3225
4.4671 | -81.3243
4.4674 | -81.324
4.4677 | -81.3232
4.468 | -81.3234
4.4684 | -81.3235
4.4687 | -81.3236
4.469 | -81.3242
4.4694 | -81.3252
4.4697 | -81.3264
4.47 | -81.3279
4.4704 | -81.3287
4.4707 | -81.3292
4.471 | -81.3308
4.4713 | -81.3324
4.4717 | -81.3332
4.472 | -81.3336
4.4723 | -81.3348
4.4727 | -81.3359
4.473 | -81.337
4.4733 | -81.3368
4.4736 | -81.3367
4.474 | -81.3374
4.4743 | -81.3386
4.4746 | -81.3391
4.475 | -81.3383
4.4753 | -81.3389
4.4756 | -81.3401
4.4759 | -81.3408
4.4763 | -81.3413
4.4766 | -81.3421
4.4769 | -81.3431
4.4773 | -81.3435
4.4776 | -81.3446
4.4779 | -81.3454
4.4783 | -81.3459
4.4786 | -81.3471
4.4789 | -81.3489
4.4792 | -81.3505
4.4796 | -81.3509
4.4799 | -81.3512
4.4802 | -81.3525
4.4806 | -81.3536
4.4809 | -81.3539
4.4812 | -81.3551
4.4815 | -81.3574
4.4819 | -81.3597
4.4822 | -81.3609
4.4825 | -81.3626
4.4829 | -81.3657
4.4832 | -81.3687
4.4835 | -81.3684
4.4838 | -81.3681
4.4842 | -81.3687
4.4845 | -81.3702
4.4848 | -81.3699
4.4852 | -81.3692
4.4855 | -81.3696
4.4858 | -81.3694
4.4862 | -81.3697
4.4865 | -81.37
4.4868 | -81.3707
4.4871 | -81.3715
4.4875 | -81.3715
4.4878 | -81.3706
4.4881 | -81.3719
4.4885 | -81.3724
4.4888 | -81.3735
4.4891 | -81.3749
4.4894 | -81.3752
4.4898 | -81.3756
4.4901 | -81.3775
4.4904 | -81.3789
4.4908 | -81.3797
4.4911 | -81.3801
4.4914 | -81.3807
4.4917 | -81.3804
4.4921 | -81.3808
4.4924 | -81.3812
4.4927 | -81.3819
4.4931 | -81.3823
4.4934 | -81.3818
4.4937 | -81.3815
4.494 | -81.3811
4.4944 | -81.3812
4.4947 | -81.3813
4.495 | -81.3824
4.4954 | -81.3827
4.4957 | -81.3834
4.496 | -81.3835
4.4964 | -81.3844
4.4967 | -81.3844
4.497 | -81.3845
4.4973 | -81.3855
4.4977 | -81.3878
4.498 | -81.3898
4.4983 | -81.3899
4.4987 | -81.3901
4.499 | -81.3913
4.4993 | -81.3931
4.4996 | -81.3933
4.5 | -81.3931
4.5003 | -81.3949
4.5006 | -81.3963
4.501 | -81.3964
4.5013 | -81.3966
4.5016 | -81.3972
4.5019 | -81.3974
4.5023 | -81.397
4.5026 | -81.3972
4.5029 | -81.3976
4.5033 | -81.3989
4.5036 | -81.4007
4.5039 | -81.403
4.5043 | -81.406
4.5046 | -81.4074
4.5049 | -81.4071
4.5052 | -81.4079
4.5056 | -81.4085
4.5059 | -81.4093
4.5062 | -81.4099
4.5066 | -81.4093
4.5069 | -81.408
4.5072 | -81.4085
4.5075 | -81.4088
4.5079 | -81.4095
4.5082 | -81.4104
4.5085 | -81.411
4.5089 | -81.4115
4.5092 | -81.412
4.5095 | -81.4137
4.5098 | -81.4162
4.5102 | -81.4172
4.5105 | -81.4171
4.5108 | -81.4175
4.5112 | -81.4189
4.5115 | -81.4198
4.5118 | -81.4201
4.5122 | -81.4207
4.5125 | -81.4217
4.5128 | -81.4219
4.5131 | -81.422
4.5135 | -81.4225
4.5138 | -81.4231
4.5141 | -81.4232
4.5145 | -81.424
4.5148 | -81.4244
4.5151 | -81.4252
4.5154 | -81.4265
4.5158 | -81.428
4.5161 | -81.4289
4.5164 | -81.4291
4.5168 | -81.4295
4.5171 | -81.4302
4.5174 | -81.4307
4.5177 | -81.4312
4.5181 | -81.4322
4.5184 | -81.4327
4.5187 | -81.4335
4.5191 | -81.4335
4.5194 | -81.4332
4.5197 | -81.4333
4.5201 | -81.4344
4.5204 | -81.4359
4.5207 | -81.4372
4.521 | -81.4377
4.5214 | -81.4379
4.5217 | -81.4381
4.522 | -81.4383
4.5224 | -81.4383
4.5227 | -81.4386
4.523 | -81.4404
4.5233 | -81.4424
4.5237 | -81.4435
4.524 | -81.4437
4.5243 | -81.4434
4.5247 | -81.4439
4.525 | -81.4441
4.5253 | -81.4435
4.5256 | -81.4431
4.526 | -81.4433
4.5263 | -81.4438
4.5266 | -81.4443
4.527 | -81.4454
4.5273 | -81.4458
4.5276 | -81.4471
4.5279 | -81.448
4.5283 | -81.4483
4.5286 | -81.4485
4.5289 | -81.449
4.5293 | -81.4486
4.5296 | -81.4489
4.5299 | -81.4492
4.5303 | -81.4496
4.5306 | -81.45
4.5309 | -81.4511
4.5312 | -81.4527
4.5316 | -81.4531
4.5319 | -81.4537
4.5322 | -81.4529
4.5326 | -81.4496
4.5329 | -81.449
4.5332 | -81.4498
4.5335 | -81.4506
4.5339 | -81.4509
4.5342 | -81.4511
4.5345 | -81.4514
4.5349 | -81.4516
4.5352 | -81.4525
4.5355 | -81.4549
4.5358 | -81.4554
4.5362 | -81.4557
4.5365 | -81.4558
4.5368 | -81.4561
4.5372 | -81.4566
4.5375 | -81.4572
4.5378 | -81.4573
4.5382 | -81.4574
4.5385 | -81.4586
4.5388 | -81.4585
4.5391 | -81.4588
4.5395 | -81.4582
4.5398 | -81.4589
4.5401 | -81.4594
4.5405 | -81.4603
4.5408 | -81.4614
4.5411 | -81.4621
4.5414 | -81.4626
4.5418 | -81.4626
4.5421 | -81.4634
4.5424 | -81.4641
4.5428 | -81.4646
4.5431 | -81.465
4.5434 | -81.4652
4.5437 | -81.4648
4.5441 | -81.4676
4.5444 | -81.4693
4.5447 | -81.4703
4.5451 | -81.4708
4.5454 | -81.4708
4.5457 | -81.4712
4.5461 | -81.4702
4.5464 | -81.4699
4.5467 | -81.4706
4.547 | -81.4709
4.5474 | -81.4708
4.5477 | -81.4708
4.548 | -81.4718
4.5484 | -81.4726
4.5487 | -81.4723
4.549 | -81.4732
4.5493 | -81.4741
4.5497 | -81.474
4.55 | -81.4738
4.5503 | -81.4741
4.5507 | -81.4749
4.551 | -81.4757
4.5513 | -81.4767
4.5516 | -81.4778
4.552 | -81.4798
4.5523 | -81.482
4.5526 | -81.4837
4.553 | -81.4841
4.5533 | -81.4838
4.5536 | -81.484
4.554 | -81.484
4.5543 | -81.483
4.5546 | -81.4828
4.5549 | -81.4828
4.5553 | -81.4821
4.5556 | -81.4823
4.5559 | -81.4832
4.5563 | -81.4845
4.5566 | -81.4851
4.5569 | -81.485
4.5572 | -81.4843
4.5576 | -81.4841
4.5579 | -81.4842
4.5582 | -81.4842
4.5586 | -81.486
4.5589 | -81.4862
4.5592 | -81.4869
4.5595 | -81.4875
4.5599 | -81.487
4.5602 | -81.487
4.5605 | -81.4871
4.5609 | -81.4854
4.5612 | -81.4848
4.5615 | -81.4851
4.5619 | -81.4861
4.5622 | -81.4862
4.5625 | -81.4869
4.5628 | -81.4879
4.5632 | -81.4885
4.5635 | -81.489
4.5638 | -81.4895
4.5642 | -81.49
4.5645 | -81.4893
4.5648 | -81.489
4.5651 | -81.4885
4.5655 | -81.488
4.5658 | -81.4874
4.5661 | -81.4886
4.5665 | -81.4892
4.5668 | -81.49
4.5671 | -81.4911
4.5674 | -81.4913
4.5678 | -81.4903
4.5681 | -81.4901
4.5684 | -81.4904
4.5688 | -81.4906
4.5691 | -81.4909
4.5694 | -81.4898
4.5698 | -81.4887
4.5701 | -81.4887
4.5704 | -81.491
4.5707 | -81.4937
4.5711 | -81.4946
4.5714 | -81.4942
4.5717 | -81.4942
4.5721 | -81.4938
4.5724 | -81.4934
4.5727 | -81.4931
4.573 | -81.4929
4.5734 | -81.4929
4.5737 | -81.4933
4.574 | -81.4936
4.5744 | -81.4944
4.5747 | -81.4947
4.575 | -81.4951
4.5753 | -81.4962
4.5757 | -81.4966
4.576 | -81.4955
4.5763 | -81.4956
4.5767 | -81.4959
4.577 | -81.4963
4.5773 | -81.4978
4.5777 | -81.5001
4.578 | -81.5007
4.5783 | -81.5016
4.5786 | -81.503
4.579 | -81.5031
4.5793 | -81.5031
4.5796 | -81.5017
4.58 | -81.5007
4.5803 | -81.5006
4.5806 | -81.5002
4.5809 | -81.5001
4.5813 | -81.5001
4.5816 | -81.5011
4.5819 | -81.5032
4.5823 | -81.5035
4.5826 | -81.5037
4.5829 | -81.5038
4.5832 | -81.5032
4.5836 | -81.5025
4.5839 | -81.5027
4.5842 | -81.5033
4.5846 | -81.5038
4.5849 | -81.5028
4.5852 | -81.5018
4.5855 | -81.5025
4.5859 | -81.5025
4.5862 | -81.5027
4.5865 | -81.5028
4.5869 | -81.5027
4.5872 | -81.5031
4.5875 | -81.502
4.5879 | -81.5002
4.5882 | -81.4999
4.5885 | -81.5011
4.5888 | -81.5022
4.5892 | -81.503
4.5895 | -81.5042
4.5898 | -81.505
4.5902 | -81.5066
4.5905 | -81.507
4.5908 | -81.5062
4.5911 | -81.5069
4.5915 | -81.5071
4.5918 | -81.5071
4.5921 | -81.5075
4.5925 | -81.5079
4.5928 | -81.5082
4.5931 | -81.5069
4.5934 | -81.5061
4.5938 | -81.5066
4.5941 | -81.5082
4.5944 | -81.5084
4.5948 | -81.508
4.5951 | -81.5082
4.5954 | -81.5085
4.5958 | -81.5088
4.5961 | -81.5099
4.5964 | -81.5094
4.5967 | -81.5077
4.5971 | -81.5069
4.5974 | -81.5066
4.5977 | -81.5064
4.5981 | -81.5073
4.5984 | -81.5067
4.5987 | -81.5066
4.599 | -81.5068
4.5994 | -81.5066
4.5997 | -81.5065
4.6 | -81.506
4.6004 | -81.5063
4.6007 | -81.5077
4.601 | -81.5086
4.6013 | -81.5088
4.6017 | -81.5093
4.602 | -81.5102
4.6023 | -81.5105
4.6027 | -81.51
4.603 | -81.5095
4.6033 | -81.5087
4.6037 | -81.5087
4.604 | -81.5096
4.6043 | -81.5106
4.6046 | -81.5105
4.605 | -81.5109
4.6053 | -81.5109
4.6056 | -81.5112
4.606 | -81.5107
4.6063 | -81.5105
4.6066 | -81.5103
4.6069 | -81.51
4.6073 | -81.5099
4.6076 | -81.51
4.6079 | -81.5096
4.6083 | -81.5097
4.6086 | -81.5098
4.6089 | -81.5111
4.6092 | -81.5117
4.6096 | -81.5115
4.6099 | -81.5119
4.6102 | -81.5121
4.6106 | -81.5122
4.6109 | -81.5127
4.6112 | -81.5133
4.6116 | -81.5149
4.6119 | -81.5156
4.6122 | -81.5158
4.6125 | -81.5161
4.6129 | -81.5158
4.6132 | -81.5157
4.6135 | -81.5159
4.6139 | -81.5164
4.6142 | -81.517
4.6145 | -81.5166
4.6148 | -81.5165
4.6152 | -81.5163
4.6155 | -81.5165
4.6158 | -81.516
4.6162 | -81.5162
4.6165 | -81.5164
4.6168 | -81.5164
4.6171 | -81.516
4.6175 | -81.5158
4.6178 | -81.5162
4.6181 | -81.5163
4.6185 | -81.517
4.6188 | -81.518
4.6191 | -81.5183
4.6195 | -81.5191
4.6198 | -81.5195
4.6201 | -81.5207
4.6204 | -81.5229
4.6208 | -81.5246
4.6211 | -81.5261
4.6214 | -81.5263
4.6218 | -81.5271
4.6221 | -81.5284
4.6224 | -81.528
4.6227 | -81.5275
4.6231 | -81.5266
4.6234 | -81.5262
4.6237 | -81.5268
4.6241 | -81.5269
4.6244 | -81.527
4.6247 | -81.5256
4.625 | -81.5259
4.6254 | -81.526
4.6257 | -81.5257
4.626 | -81.5258
4.6264 | -81.5263
4.6267 | -81.5265
4.627 | -81.5286
4.6273 | -81.5297
4.6277 | -81.5294
4.628 | -81.5299
4.6283 | -81.5302
4.6287 | -81.5303
4.629 | -81.5302
4.6293 | -81.5316
4.6297 | -81.5346
4.63 | -81.5359
4.6303 | -81.5351
4.6306 | -81.5352
4.631 | -81.5347
4.6313 | -81.5323
4.6316 | -81.5311
4.632 | -81.5308
4.6323 | -81.531
4.6326 | -81.5309
4.6329 | -81.5303
4.6333 | -81.5297
4.6336 | -81.529
4.6339 | -81.5297
4.6343 | -81.5301
4.6346 | -81.5301
4.6349 | -81.5307
4.6352 | -81.5317
4.6356 | -81.5328
4.6359 | -81.5324
4.6362 | -81.5321
4.6366 | -81.5317
4.6369 | -81.5312
4.6372 | -81.5307
4.6376 | -81.5298
4.6379 | -81.5293
4.6382 | -81.5294
4.6385 | -81.53
4.6389 | -81.5315
4.6392 | -81.5324
4.6395 | -81.5327
4.6399 | -81.5322
4.6402 | -81.5311
4.6405 | -81.5312
4.6408 | -81.5319
4.6412 | -81.5329
4.6415 | -81.5332
4.6418 | -81.5331
4.6422 | -81.5328
4.6425 | -81.5325
4.6428 | -81.5315
4.6431 | -81.5309
4.6435 | -81.5315
4.6438 | -81.532
4.6441 | -81.534
4.6445 | -81.534
4.6448 | -81.5345
4.6451 | -81.5348
4.6455 | -81.5346
4.6458 | -81.5336
4.6461 | -81.5328
4.6464 | -81.5325
4.6468 | -81.5324
4.6471 | -81.5315
4.6474 | -81.5318
4.6478 | -81.532
4.6481 | -81.5319
4.6484 | -81.5318
4.6487 | -81.532
4.6491 | -81.5316
4.6494 | -81.531
4.6497 | -81.5303
4.6501 | -81.5303
4.6504 | -81.532
4.6507 | -81.5319
4.651 | -81.5321
4.6514 | -81.5319
4.6517 | -81.5328
4.652 | -81.5328
4.6524 | -81.5328
4.6527 | -81.5328
4.653 | -81.5325
4.6534 | -81.5309
4.6537 | -81.5287
4.654 | -81.5291
4.6543 | -81.5293
4.6547 | -81.5292
4.655 | -81.5295
4.6553 | -81.5303
4.6557 | -81.5308
4.656 | -81.5314
4.6563 | -81.5315
4.6566 | -81.5313
4.657 | -81.5312
4.6573 | -81.5316
4.6576 | -81.5318
4.658 | -81.5319
4.6583 | -81.5313
4.6586 | -81.5307
4.6589 | -81.5304
4.6593 | -81.5298
4.6596 | -81.5298
4.6599 | -81.5294
4.6603 | -81.5301
4.6606 | -81.5304
4.6609 | -81.5307
4.6613 | -81.5311
4.6616 | -81.5314
4.6619 | -81.5314
4.6622 | -81.5318
4.6626 | -81.5322
4.6629 | -81.5337
4.6632 | -81.5338
4.6636 | -81.5327
4.6639 | -81.5317
4.6642 | -81.5308
4.6645 | -81.5325
4.6649 | -81.5342
4.6652 | -81.5343
4.6655 | -81.534
4.6659 | -81.5341
4.6662 | -81.5336
4.6665 | -81.5328
4.6668 | -81.5324
4.6672 | -81.5328
4.6675 | -81.5328
4.6678 | -81.5308
4.6682 | -81.5294
4.6685 | -81.5286
4.6688 | -81.5272
4.6692 | -81.5262
4.6695 | -81.5263
4.6698 | -81.5265
4.6701 | -81.5269
4.6705 | -81.5268
4.6708 | -81.5265
4.6711 | -81.5266
4.6715 | -81.5278
4.6718 | -81.5287
4.6721 | -81.5296
4.6724 | -81.5295
4.6728 | -81.5301
4.6731 | -81.5303
4.6734 | -81.53
4.6738 | -81.5306
4.6741 | -81.5324
4.6744 | -81.5332
4.6747 | -81.5327
4.6751 | -81.5325
4.6754 | -81.5326
4.6757 | -81.5325
4.6761 | -81.5326
4.6764 | -81.533
4.6767 | -81.5329
4.6771 | -81.5329
4.6774 | -81.5329
4.6777 | -81.5321
4.678 | -81.5324
4.6784 | -81.5325
4.6787 | -81.5318
4.679 | -81.5319
4.6794 | -81.5319
4.6797 | -81.5315
4.68 | -81.5317
4.6803 | -81.5324
4.6807 | -81.5329
4.681 | -81.5337
4.6813 | -81.5336
4.6817 | -81.5341
4.682 | -81.5341
4.6823 | -81.5348
4.6826 | -81.5352
4.683 | -81.5351
4.6833 | -81.5351
4.6836 | -81.5345
4.684 | -81.5318
4.6843 | -81.5293
4.6846 | -81.5287
4.6849 | -81.5286
4.6853 | -81.5283
4.6856 | -81.5281
4.6859 | -81.5283
4.6863 | -81.529
4.6866 | -81.5294
4.6869 | -81.529
4.6873 | -81.5284
4.6876 | -81.5286
4.6879 | -81.529
4.6882 | -81.5299
4.6886 | -81.5306
4.6889 | -81.5298
4.6892 | -81.5297
4.6896 | -81.5299
4.6899 | -81.5296
4.6902 | -81.5286
4.6905 | -81.5272
4.6909 | -81.5271
4.6912 | -81.5277
4.6915 | -81.5276
4.6919 | -81.5283
4.6922 | -81.5283
4.6925 | -81.5287
4.6928 | -81.5291
4.6932 | -81.5305
4.6935 | -81.5312
4.6938 | -81.5311
4.6942 | -81.5316
4.6945 | -81.5323
4.6948 | -81.5331
4.6952 | -81.5334
4.6955 | -81.5334
4.6958 | -81.5334
4.6961 | -81.5335
4.6965 | -81.5331
4.6968 | -81.5335
4.6971 | -81.5338
4.6975 | -81.5338
4.6978 | -81.5335
4.6981 | -81.5339
4.6984 | -81.5347
4.6988 | -81.5349
4.6991 | -81.535
4.6994 | -81.5356
4.6998 | -81.5363
4.7001 | -81.5366
4.7004 | -81.537
4.7007 | -81.537
4.7011 | -81.5376
4.7014 | -81.5379
4.7017 | -81.5381
4.7021 | -81.5379
4.7024 | -81.5389
4.7027 | -81.539
4.7031 | -81.5392
4.7034 | -81.5398
4.7037 | -81.5407
4.704 | -81.5406
4.7044 | -81.54
4.7047 | -81.5389
4.705 | -81.5388
4.7054 | -81.5387
4.7057 | -81.5388
4.706 | -81.5393
4.7063 | -81.5395
4.7067 | -81.5407
4.707 | -81.5408
4.7073 | -81.54
4.7077 | -81.5413
4.708 | -81.5418
4.7083 | -81.542
4.7086 | -81.5422
4.709 | -81.5424
4.7093 | -81.5429
4.7096 | -81.5435
4.71 | -81.5445
4.7103 | -81.5451
4.7106 | -81.5459
4.711 | -81.5457
4.7113 | -81.546
4.7116 | -81.5462
4.7119 | -81.5452
4.7123 | -81.5436
4.7126 | -81.5435
4.7129 | -81.5438
4.7133 | -81.544
4.7136 | -81.5441
4.7139 | -81.5447
4.7142 | -81.5451
4.7146 | -81.5437
4.7149 | -81.5425
4.7152 | -81.5437
4.7156 | -81.5461
4.7159 | -81.5473
4.7162 | -81.5485
4.7165 | -81.5493
4.7169 | -81.5506
4.7172 | -81.5526
4.7175 | -81.5536
4.7179 | -81.5536
4.7182 | -81.5529
4.7185 | -81.552
4.7188 | -81.5518
4.7192 | -81.5518
4.7195 | -81.5516
4.7198 | -81.5513
4.7202 | -81.5523
4.7205 | -81.5523
4.7208 | -81.553
4.7212 | -81.5524
4.7215 | -81.5521
4.7218 | -81.552
4.7221 | -81.552
4.7225 | -81.5521
4.7228 | -81.5523
4.7231 | -81.5532
4.7235 | -81.553
4.7238 | -81.5527
4.7241 | -81.5524
4.7244 | -81.5516
4.7248 | -81.5516
4.7251 | -81.5532
4.7254 | -81.5544
4.7258 | -81.5538
4.7261 | -81.5532
4.7264 | -81.5535
4.7268 | -81.5532
4.7271 | -81.552
4.7274 | -81.5519
4.7277 | -81.5522
4.7281 | -81.5533
4.7284 | -81.5536
4.7287 | -81.5539
4.7291 | -81.5554
4.7294 | -81.5554
4.7297 | -81.5555
4.73 | -81.5565
4.7304 | -81.5579
4.7307 | -81.5592
4.731 | -81.56
4.7314 | -81.5603
4.7317 | -81.5613
4.732 | -81.5618
4.7323 | -81.5615
4.7327 | -81.5606
4.733 | -81.5608
4.7333 | -81.5618
4.7337 | -81.5633
4.734 | -81.5629
4.7343 | -81.5643
4.7346 | -81.5648
4.735 | -81.5652
4.7353 | -81.5646
4.7356 | -81.5645
4.736 | -81.5644
4.7363 | -81.5644
4.7366 | -81.5647
4.737 | -81.565
4.7373 | -81.5653
4.7376 | -81.5657
4.7379 | -81.5663
4.7383 | -81.5666
4.7386 | -81.5682
4.7389 | -81.5686
4.7393 | -81.5689
4.7396 | -81.569
4.7399 | -81.5704
4.7402 | -81.5723
4.7406 | -81.573
4.7409 | -81.5736
4.7412 | -81.5751
4.7416 | -81.5761
4.7419 | -81.5758
4.7422 | -81.5756
4.7425 | -81.576
4.7429 | -81.5764
4.7432 | -81.5768
4.7435 | -81.5767
4.7439 | -81.5771
4.7442 | -81.5777
4.7445 | -81.5789
4.7449 | -81.5795
4.7452 | -81.5803
4.7455 | -81.5809
4.7458 | -81.5814
4.7462 | -81.5817
4.7465 | -81.5831
4.7468 | -81.5843
4.7472 | -81.586
4.7475 | -81.5867
4.7478 | -81.5868
4.7481 | -81.5883
4.7485 | -81.5885
4.7488 | -81.5883
4.7491 | -81.5881
4.7495 | -81.5883
4.7498 | -81.5897
4.7501 | -81.5905
4.7504 | -81.5912
4.7508 | -81.5919
4.7511 | -81.592
4.7514 | -81.593
4.7518 | -81.5931
4.7521 | -81.5937
4.7524 | -81.5943
4.7528 | -81.5948
4.7531 | -81.5965
4.7534 | -81.5996
4.7537 | -81.6012
4.7541 | -81.6017
4.7544 | -81.6018
4.7547 | -81.6014
4.7551 | -81.6021
4.7554 | -81.603
4.7557 | -81.6041
4.756 | -81.6043
4.7564 | -81.6032
4.7567 | -81.6038
4.757 | -81.6044
4.7574 | -81.6048
4.7577 | -81.6057
4.758 | -81.6061
4.7583 | -81.6071
4.7587 | -81.6078
4.759 | -81.6085
4.7593 | -81.609
4.7597 | -81.6098
4.76 | -81.611
4.7603 | -81.6122
4.7607 | -81.6119
4.761 | -81.6119
4.7613 | -81.6114
4.7616 | -81.6122
4.762 | -81.6129
4.7623 | -81.6133
4.7626 | -81.6144
4.763 | -81.6173
4.7633 | -81.6217
4.7636 | -81.6242
4.7639 | -81.6253
4.7643 | -81.627
4.7646 | -81.6277
4.7649 | -81.629
4.7653 | -81.6305
4.7656 | -81.6316
4.7659 | -81.6324
4.7662 | -81.6329
4.7666 | -81.634
4.7669 | -81.6341
4.7672 | -81.634
4.7676 | -81.6344
4.7679 | -81.635
4.7682 | -81.6359
4.7686 | -81.6354
4.7689 | -81.6367
4.7692 | -81.6398
4.7695 | -81.641
4.7699 | -81.6417
4.7702 | -81.6425
4.7705 | -81.6433
4.7709 | -81.6435
4.7712 | -81.6441
4.7715 | -81.6458
4.7718 | -81.6478
4.7722 | -81.6478
4.7725 | -81.6485
4.7728 | -81.6493
4.7732 | -81.6472
4.7735 | -81.6446
4.7738 | -81.6447
4.7741 | -81.6456
4.7745 | -81.6466
4.7748 | -81.6473
4.7751 | -81.6479
4.7755 | -81.649
4.7758 | -81.65
4.7761 | -81.6508
4.7764 | -81.6521
4.7768 | -81.6529
4.7771 | -81.6553
4.7774 | -81.6555
4.7778 | -81.6548
4.7781 | -81.655
4.7784 | -81.6547
4.7788 | -81.6549
4.7791 | -81.6557
4.7794 | -81.6554
4.7797 | -81.6568
4.7801 | -81.6574
4.7804 | -81.6575
4.7807 | -81.658
4.7811 | -81.6582
4.7814 | -81.659
4.7817 | -81.66
4.782 | -81.6607
4.7824 | -81.6615
4.7827 | -81.6618
4.783 | -81.6627
4.7834 | -81.6635
4.7837 | -81.6641
4.784 | -81.6646
4.7843 | -81.6646
4.7847 | -81.663
4.785 | -81.6628
4.7853 | -81.6629
4.7857 | -81.663
4.786 | -81.6646
4.7863 | -81.6652
4.7867 | -81.6663
4.787 | -81.6693
4.7873 | -81.6705
4.7876 | -81.671
4.788 | -81.6718
4.7883 | -81.6724
4.7886 | -81.6736
4.789 | -81.6765
4.7893 | -81.6789
4.7896 | -81.6802
4.7899 | -81.682
4.7903 | -81.6838
4.7906 | -81.6843
4.7909 | -81.6851
4.7913 | -81.6857
4.7916 | -81.6854
4.7919 | -81.6851
4.7922 | -81.6857
4.7926 | -81.687
4.7929 | -81.6875
4.7932 | -81.6877
4.7936 | -81.6885
4.7939 | -81.6896
4.7942 | -81.6908
4.7946 | -81.6917
4.7949 | -81.6927
4.7952 | -81.6947
4.7955 | -81.6966
4.7959 | -81.6983
4.7962 | -81.6995
4.7965 | -81.7001
4.7969 | -81.7
4.7972 | -81.7016
4.7975 | -81.7025
4.7978 | -81.7035
4.7982 | -81.7047
4.7985 | -81.7056
4.7988 | -81.7065
4.7992 | -81.7058
4.7995 | -81.7058
4.7998 | -81.7068
4.8001 | -81.7079
4.8005 | -81.7086
4.8008 | -81.7094
4.8011 | -81.71
4.8015 | -81.711
4.8018 | -81.7129
4.8021 | -81.7144
4.8025 | -81.7152
4.8028 | -81.716
4.8031 | -81.7173
4.8034 | -81.718
4.8038 | -81.7178
4.8041 | -81.7185
4.8044 | -81.7194
4.8048 | -81.7198
4.8051 | -81.7207
4.8054 | -81.7211
4.8057 | -81.7207
4.8061 | -81.721
4.8064 | -81.7224
4.8067 | -81.7225
4.8071 | -81.7222
4.8074 | -81.723
4.8077 | -81.7248
4.808 | -81.7267
4.8084 | -81.7277
4.8087 | -81.7293
4.809 | -81.7304
4.8094 | -81.7316
4.8097 | -81.7322
4.81 | -81.7332
4.8104 | -81.735
4.8107 | -81.7369
4.811 | -81.7376
4.8113 | -81.738
4.8117 | -81.7388
4.812 | -81.7393
4.8123 | -81.7408
4.8127 | -81.7422
4.813 | -81.7431
4.8133 | -81.7434
4.8136 | -81.7441
4.814 | -81.7451
4.8143 | -81.7458
4.8146 | -81.7472
4.815 | -81.7481
4.8153 | -81.75
4.8156 | -81.7513
4.8159 | -81.7522
4.8163 | -81.753
4.8166 | -81.7539
4.8169 | -81.7558
4.8173 | -81.7575
4.8176 | -81.7593
4.8179 | -81.7607
4.8182 | -81.762
4.8186 | -81.763
4.8189 | -81.7641
4.8192 | -81.7646
4.8196 | -81.7652
4.8199 | -81.7664
4.8202 | -81.7681
4.8206 | -81.7692
4.8209 | -81.7701
4.8212 | -81.7706
4.8215 | -81.7713
4.8219 | -81.7726
4.8222 | -81.7747
4.8225 | -81.7761
4.8229 | -81.7773
4.8232 | -81.7786
4.8235 | -81.7799
4.8238 | -81.7805
4.8242 | -81.7817
4.8245 | -81.7822
4.8248 | -81.7824
4.8252 | -81.7828
4.8255 | -81.7841
4.8258 | -81.7845
4.8261 | -81.7857
4.8265 | -81.7874
4.8268 | -81.7884
4.8271 | -81.7892
4.8275 | -81.7891
4.8278 | -81.7889
4.8281 | -81.7901
4.8285 | -81.7901
4.8288 | -81.7909
4.8291 | -81.7929
4.8294 | -81.7942
4.8298 | -81.795
4.8301 | -81.796
4.8304 | -81.7969
4.8308 | -81.7988
4.8311 | -81.8015
4.8314 | -81.8014
4.8317 | -81.8017
4.8321 | -81.8025
4.8324 | -81.8033
4.8327 | -81.8036
4.8331 | -81.804
4.8334 | -81.805
4.8337 | -81.8056
4.834 | -81.8063
4.8344 | -81.806
4.8347 | -81.8053
4.835 | -81.8054
4.8354 | -81.8075
4.8357 | -81.8083
4.836 | -81.8074
4.8364 | -81.8061
4.8367 | -81.8062
4.837 | -81.8068
4.8373 | -81.8075
4.8377 | -81.8086
4.838 | -81.8101
4.8383 | -81.8112
4.8387 | -81.8118
4.839 | -81.8129
4.8393 | -81.8137
4.8396 | -81.8145
4.84 | -81.8145
4.8403 | -81.8154
4.8406 | -81.8159
4.841 | -81.8156
4.8413 | -81.8166
4.8416 | -81.8173
4.8419 | -81.8178
4.8423 | -81.8182
4.8426 | -81.8188
4.8429 | -81.8199
4.8433 | -81.8207
4.8436 | -81.8215
4.8439 | -81.8223
4.8443 | -81.8233
4.8446 | -81.8245
4.8449 | -81.8261
4.8452 | -81.8268
4.8456 | -81.8269
4.8459 | -81.8276
4.8462 | -81.8284
4.8466 | -81.8292
4.8469 | -81.8304
4.8472 | -81.8315
4.8475 | -81.8321
4.8479 | -81.8324
4.8482 | -81.833
4.8485 | -81.8334
4.8489 | -81.8341
4.8492 | -81.8346
4.8495 | -81.8346
4.8498 | -81.8359
4.8502 | -81.8363
4.8505 | -81.8363
4.8508 | -81.8379
4.8512 | -81.8398
4.8515 | -81.8412
4.8518 | -81.8417
4.8522 | -81.8426
4.8525 | -81.8437
4.8528 | -81.8448
4.8531 | -81.8452
4.8535 | -81.8441
4.8538 | -81.8449
4.8541 | -81.8458
4.8545 | -81.8458
4.8548 | -81.846
4.8551 | -81.8468
4.8554 | -81.8476
4.8558 | -81.8485
4.8561 | -81.8493
4.8564 | -81.8496
4.8568 | -81.8503
4.8571 | -81.8511
4.8574 | -81.8519
4.8577 | -81.8526
4.8581 | -81.8532
4.8584 | -81.8548
4.8587 | -81.8586
4.8591 | -81.8601
4.8594 | -81.8606
4.8597 | -81.8608
4.8601 | -81.8613
4.8604 | -81.8621
4.8607 | -81.8621
4.861 | -81.8626
4.8614 | -81.8635
4.8617 | -81.8626
4.862 | -81.8613
4.8624 | -81.8614
4.8627 | -81.8622
4.863 | -81.8631
4.8633 | -81.8645
4.8637 | -81.8654
4.864 | -81.8648
4.8643 | -81.8645
4.8647 | -81.8666
4.865 | -81.8677
4.8653 | -81.8692
4.8656 | -81.8696
4.866 | -81.8703
4.8663 | -81.8719
4.8666 | -81.8722
4.867 | -81.8725
4.8673 | -81.8727
4.8676 | -81.8731
4.868 | -81.8737
4.8683 | -81.8736
4.8686 | -81.8742
4.8689 | -81.8746
4.8693 | -81.8746
4.8696 | -81.8746
4.8699 | -81.8749
4.8703 | -81.8756
4.8706 | -81.8772
4.8709 | -81.8772
4.8712 | -81.8774
4.8716 | -81.8783
4.8719 | -81.8787
4.8722 | -81.8794
4.8726 | -81.8793
4.8729 | -81.8794
4.8732 | -81.8814
4.8735 | -81.8827
4.8739 | -81.8828
4.8742 | -81.8823
4.8745 | -81.8819
4.8749 | -81.8829
4.8752 | -81.8837
4.8755 | -81.884
4.8758 | -81.8849
4.8762 | -81.8861
4.8765 | -81.8864
4.8768 | -81.8865
4.8772 | -81.8872
4.8775 | -81.886
4.8778 | -81.885
4.8782 | -81.8854
4.8785 | -81.8854
4.8788 | -81.885
4.8791 | -81.8854
4.8795 | -81.8851
4.8798 | -81.8851
4.8801 | -81.885
4.8805 | -81.8854
4.8808 | -81.8863
4.8811 | -81.8863
4.8814 | -81.886
4.8818 | -81.8861
4.8821 | -81.8876
4.8824 | -81.888
4.8828 | -81.8886
4.8831 | -81.8898
4.8834 | -81.8903
4.8837 | -81.8897
4.8841 | -81.8895
4.8844 | -81.8897
4.8847 | -81.8903
4.8851 | -81.892
4.8854 | -81.8929
4.8857 | -81.8929
4.8861 | -81.8931
4.8864 | -81.8927
4.8867 | -81.8923
4.887 | -81.893
4.8874 | -81.8931
4.8877 | -81.8911
4.888 | -81.8898
4.8884 | -81.89
4.8887 | -81.8896
4.889 | -81.8912
4.8893 | -81.8923
4.8897 | -81.8924
4.89 | -81.8923
4.8903 | -81.8924
4.8907 | -81.8913
4.891 | -81.8917
4.8913 | -81.8923
4.8916 | -81.8931
4.892 | -81.8932
4.8923 | -81.8925
4.8926 | -81.8922
4.893 | -81.893
4.8933 | -81.894
4.8936 | -81.8953
4.894 | -81.8961
4.8943 | -81.8974
4.8946 | -81.8976
4.8949 | -81.8968
4.8953 | -81.8962
4.8956 | -81.8963
4.8959 | -81.8964
4.8963 | -81.897
4.8966 | -81.8983
4.8969 | -81.8991
4.8972 | -81.8995
4.8976 | -81.8996
4.8979 | -81.8998
4.8982 | -81.9014
4.8986 | -81.9023
4.8989 | -81.9019
4.8992 | -81.9009
4.8995 | -81.9002
4.8999 | -81.9004
4.9002 | -81.9005
4.9005 | -81.9012
4.9009 | -81.9022
4.9012 | -81.9022
4.9015 | -81.9014
4.9019 | -81.902
4.9022 | -81.9016
4.9025 | -81.9016
4.9028 | -81.9012
4.9032 | -81.9009
4.9035 | -81.9009
4.9038 | -81.9009
4.9042 | -81.9008
4.9045 | -81.9007
4.9048 | -81.9005
4.9051 | -81.9009
4.9055 | -81.9001
4.9058 | -81.9005
4.9061 | -81.9007
4.9065 | -81.901
4.9068 | -81.9001
4.9071 | -81.8978
4.9074 | -81.8973
4.9078 | -81.8975
4.9081 | -81.8978
4.9084 | -81.898
4.9088 | -81.8977
4.9091 | -81.8985
4.9094 | -81.8978
4.9097 | -81.8991
4.9101 | -81.8993
4.9104 | -81.8996
4.9107 | -81.8985
4.9111 | -81.8975
4.9114 | -81.8974
4.9117 | -81.8979
4.9121 | -81.8984
4.9124 | -81.8983
4.9127 | -81.8984
4.913 | -81.8993
4.9134 | -81.9005
4.9137 | -81.9006
4.914 | -81.9008
4.9144 | -81.901
4.9147 | -81.9023
4.915 | -81.9035
4.9153 | -81.9035
4.9157 | -81.9037
4.916 | -81.9036
4.9163 | -81.9038
4.9167 | -81.9044
4.917 | -81.9048
4.9173 | -81.9047
4.9177 | -81.9057
4.918 | -81.9074
4.9183 | -81.9076
4.9186 | -81.9079
4.919 | -81.9085
4.9193 | -81.9085
4.9196 | -81.9083
4.92 | -81.9083
4.9203 | -81.9084
4.9206 | -81.9077
4.9209 | -81.907
4.9213 | -81.9076
4.9216 | -81.9077
4.9219 | -81.9077
4.9223 | -81.9063
4.9226 | -81.9055
4.9229 | -81.9047
4.9232 | -81.9044
4.9236 | -81.9058
4.9239 | -81.9063
4.9242 | -81.9065
4.9246 | -81.9065
4.9249 | -81.9063
4.9252 | -81.9062
4.9255 | -81.9055
4.9259 | -81.9052
4.9262 | -81.905
4.9265 | -81.9043
4.9269 | -81.9031
4.9272 | -81.9028
4.9275 | -81.9046
4.9279 | -81.9045
4.9282 | -81.9034
4.9285 | -81.9039
4.9288 | -81.9041
4.9292 | -81.9035
4.9295 | -81.9037
4.9298 | -81.9031
4.9302 | -81.9026
4.9305 | -81.903
4.9308 | -81.9035
4.9311 | -81.9038
4.9315 | -81.904
4.9318 | -81.904
4.9321 | -81.9043
4.9325 | -81.9044
4.9328 | -81.9043
4.9331 | -81.9044
4.9334 | -81.9046
4.9338 | -81.9049
4.9341 | -81.9052
4.9344 | -81.9059
4.9348 | -81.9063
4.9351 | -81.9069
4.9354 | -81.9072
4.9358 | -81.9076
4.9361 | -81.9074
4.9364 | -81.9075
4.9367 | -81.9073
4.9371 | -81.9079
4.9374 | -81.908
4.9377 | -81.9079
4.9381 | -81.9085
4.9384 | -81.9093
4.9387 | -81.91
4.939 | -81.9097
4.9394 | -81.9096
4.9397 | -81.9103
4.94 | -81.9101
4.9404 | -81.9102
4.9407 | -81.9098
4.941 | -81.9091
4.9413 | -81.9082
4.9417 | -81.9077
4.942 | -81.9077
4.9423 | -81.9075
4.9427 | -81.9073
4.943 | -81.9065
4.9433 | -81.9059
4.9437 | -81.9061
4.944 | -81.9058
4.9443 | -81.9051
4.9446 | -81.905
4.945 | -81.905
4.9453 | -81.9044
4.9456 | -81.903
4.946 | -81.9021
4.9463 | -81.9026
4.9466 | -81.9028
4.9469 | -81.9032
4.9473 | -81.9038
4.9476 | -81.9038
4.9479 | -81.9036
4.9483 | -81.9042
4.9486 | -81.9037
4.9489 | -81.9037
4.9492 | -81.9037
4.9496 | -81.9041
4.9499 | -81.9043
4.9502 | -81.9039
4.9506 | -81.9041
4.9509 | -81.9032
4.9512 | -81.9019
4.9516 | -81.9019
4.9519 | -81.9012
4.9522 | -81.9006
4.9525 | -81.9008
4.9529 | -81.9009
4.9532 | -81.9009
4.9535 | -81.9006
4.9539 | -81.8998
4.9542 | -81.8998
4.9545 | -81.9
4.9548 | -81.9004
4.9552 | -81.9012
4.9555 | -81.9023
4.9558 | -81.9032
4.9562 | -81.9031
4.9565 | -81.9036
4.9568 | -81.9045
4.9571 | -81.9044
4.9575 | -81.9055
4.9578 | -81.9059
4.9581 | -81.9055
4.9585 | -81.9052
4.9588 | -81.9057
4.9591 | -81.9054
4.9595 | -81.9055
4.9598 | -81.9054
4.9601 | -81.9052
4.9604 | -81.9052
4.9608 | -81.9057
4.9611 | -81.9059
4.9614 | -81.9067
4.9618 | -81.9072
4.9621 | -81.907
4.9624 | -81.9067
4.9627 | -81.9066
4.9631 | -81.9066
4.9634 | -81.9081
4.9637 | -81.9085
4.9641 | -81.9095
4.9644 | -81.912
4.9647 | -81.9137
4.965 | -81.9142
4.9654 | -81.9142
4.9657 | -81.9142
4.966 | -81.9143
4.9664 | -81.9147
4.9667 | -81.9151
4.967 | -81.916
4.9673 | -81.9164
4.9677 | -81.9159
4.968 | -81.9157
4.9683 | -81.916
4.9687 | -81.9167
4.969 | -81.9175
4.9693 | -81.9177
4.9697 | -81.9185
4.97 | -81.9192
4.9703 | -81.9192
4.9706 | -81.92
4.971 | -81.9209
4.9713 | -81.9215
4.9716 | -81.922
4.972 | -81.9216
4.9723 | -81.9217
4.9726 | -81.9218
4.9729 | -81.9217
4.9733 | -81.9215
4.9736 | -81.9224
4.9739 | -81.9233
4.9743 | -81.9232
4.9746 | -81.9233
4.9749 | -81.9233
4.9752 | -81.9231
4.9756 | -81.9234
4.9759 | -81.9241
4.9762 | -81.9243
4.9766 | -81.9245
4.9769 | -81.9253
4.9772 | -81.9259
4.9776 | -81.9257
4.9779 | -81.9256
4.9782 | -81.926
4.9785 | -81.9262
4.9789 | -81.9263
4.9792 | -81.9266
4.9795 | -81.9266
4.9799 | -81.9266
4.9802 | -81.9278
4.9805 | -81.9294
4.9808 | -81.9303
4.9812 | -81.9306
4.9815 | -81.9311
4.9818 | -81.9308
4.9822 | -81.9303
4.9825 | -81.9304
4.9828 | -81.9304
4.9831 | -81.9289
4.9835 | -81.9273
4.9838 | -81.9281
4.9841 | -81.9286
4.9845 | -81.929
4.9848 | -81.9294
4.9851 | -81.9297
4.9855 | -81.9298
4.9858 | -81.9294
4.9861 | -81.9294
4.9864 | -81.9298
4.9868 | -81.9301
4.9871 | -81.9299
4.9874 | -81.9301
4.9878 | -81.9304
4.9881 | -81.9314
4.9884 | -81.933
4.9887 | -81.9356
4.9891 | -81.9367
4.9894 | -81.937
4.9897 | -81.9374
4.9901 | -81.9376
4.9904 | -81.9377
4.9907 | -81.9378
4.991 | -81.9381
4.9914 | -81.9381
4.9917 | -81.9378
4.992 | -81.9368
4.9924 | -81.9364
4.9927 | -81.9361
4.993 | -81.9366
4.9934 | -81.9374
4.9937 | -81.9377
4.994 | -81.9384
4.9943 | -81.9394
4.9947 | -81.9394
4.995 | -81.9402
4.9953 | -81.941
4.9957 | -81.9413
4.996 | -81.9419
4.9963 | -81.942
4.9966 | -81.9413
4.997 | -81.9413
4.9973 | -81.9416
4.9976 | -81.9417
4.998 | -81.9417
4.9983 | -81.9402
4.9986 | -81.9394
4.9989 | -81.9394
4.9993 | -81.9403
4.9996 | -81.9424
4.9999 | -81.9439
5.0003 | -81.9446
5.0006 | -81.9446
5.0009 | -81.9441
5.0013 | -81.9442
5.0016 | -81.9448
5.0019 | -81.9456
5.0022 | -81.9472
5.0026 | -81.9479
5.0029 | -81.9475
5.0032 | -81.948
5.0036 | -81.9505
5.0039 | -81.9522
5.0042 | -81.9532
5.0045 | -81.9541
5.0049 | -81.9562
5.0052 | -81.9574
5.0055 | -81.9578
5.0059 | -81.958
5.0062 | -81.9587
5.0065 | -81.9595
5.0068 | -81.9594
5.0072 | -81.9598
5.0075 | -81.9603
5.0078 | -81.9605
5.0082 | -81.9604
5.0085 | -81.9594
5.0088 | -81.9594
5.0091 | -81.9597
5.0095 | -81.9606
5.0098 | -81.9613
5.0101 | -81.9615
5.0105 | -81.9622
5.0108 | -81.9627
5.0111 | -81.9643
5.0115 | -81.9653
5.0118 | -81.9648
5.0121 | -81.965
5.0124 | -81.9653
5.0128 | -81.9661
5.0131 | -81.9672
5.0134 | -81.968
5.0138 | -81.9683
5.0141 | -81.969
5.0144 | -81.9703
5.0147 | -81.9708
5.0151 | -81.9719
5.0154 | -81.9728
5.0157 | -81.9736
5.0161 | -81.9743
5.0164 | -81.9753
5.0167 | -81.9759
5.017 | -81.9764
5.0174 | -81.9772
5.0177 | -81.9777
5.018 | -81.9782
5.0184 | -81.979
5.0187 | -81.9797
5.019 | -81.9804
5.0194 | -81.9811
5.0197 | -81.9819
5.02 | -81.9829
5.0203 | -81.9837
5.0207 | -81.9851
5.021 | -81.9892
5.0213 | -81.9926
5.0217 | -81.9936
5.022 | -81.9943
5.0223 | -81.9953
5.0226 | -81.996
5.023 | -81.9966
5.0233 | -81.9964
5.0236 | -81.996
5.024 | -81.9973
5.0243 | -81.9985
5.0246 | -81.9994
5.0249 | -82.0001
5.0253 | -82.0009
5.0256 | -82.0018
5.0259 | -82.0026
5.0263 | -82.0043
5.0266 | -82.0051
5.0269 | -82.0059
5.0273 | -82.0068
5.0276 | -82.0085
5.0279 | -82.0109
5.0282 | -82.0121
5.0286 | -82.0129
5.0289 | -82.0138
5.0292 | -82.0149
5.0296 | -82.0158
5.0299 | -82.016
5.0302 | -82.0154
5.0305 | -82.0153
5.0309 | -82.0154
5.0312 | -82.0164
5.0315 | -82.0169
5.0319 | -82.0161
5.0322 | -82.0169
5.0325 | -82.0174
5.0328 | -82.0194
5.0332 | -82.0203
5.0335 | -82.0213
5.0338 | -82.0221
5.0342 | -82.0231
5.0345 | -82.0244
5.0348 | -82.0257
5.0352 | -82.0268
5.0355 | -82.0281
5.0358 | -82.0294
5.0361 | -82.0312
5.0365 | -82.0324
5.0368 | -82.0332
5.0371 | -82.0342
5.0375 | -82.0349
5.0378 | -82.0354
5.0381 | -82.0367
5.0384 | -82.0384
5.0388 | -82.0396
5.0391 | -82.0403
5.0394 | -82.0407
5.0398 | -82.0416
5.0401 | -82.0418
5.0404 | -82.0423
5.0407 | -82.0435
5.0411 | -82.0447
5.0414 | -82.0426
5.0417 | -82.0428
5.0421 | -82.0438
5.0424 | -82.046
5.0427 | -82.0476
5.0431 | -82.0501
5.0434 | -82.053
5.0437 | -82.054
5.044 | -82.0547
5.0444 | -82.056
5.0447 | -82.0572
5.045 | -82.0582
5.0454 | -82.0594
5.0457 | -82.0612
5.046 | -82.0633
5.0463 | -82.064
5.0467 | -82.0655
5.047 | -82.0664
5.0473 | -82.0679
5.0477 | -82.0694
5.048 | -82.0706
5.0483 | -82.0719
5.0486 | -82.0725
5.049 | -82.0729
5.0493 | -82.0745
5.0496 | -82.0755
5.05 | -82.0761
5.0503 | -82.0759
5.0506 | -82.0759
5.051 | -82.078
5.0513 | -82.0808
5.0516 | -82.0822
5.0519 | -82.0844
5.0523 | -82.0856
5.0526 | -82.0868
5.0529 | -82.088
5.0533 | -82.0897
5.0536 | -82.0915
5.0539 | -82.0945
5.0542 | -82.0973
5.0546 | -82.0993
5.0549 | -82.0994
5.0552 | -82.101
5.0556 | -82.1028
5.0559 | -82.104
5.0562 | -82.1052
5.0565 | -82.1057
5.0569 | -82.106
5.0572 | -82.1067
5.0575 | -82.1079
5.0579 | -82.1097
5.0582 | -82.1112
5.0585 | -82.1126
5.0588 | -82.1143
5.0592 | -82.1168
5.0595 | -82.119
5.0598 | -82.1203
5.0602 | -82.1216
5.0605 | -82.122
5.0608 | -82.1227
5.0612 | -82.1244
5.0615 | -82.1266
5.0618 | -82.1281
5.0621 | -82.1296
5.0625 | -82.1305
5.0628 | -82.1315
5.0631 | -82.1328
5.0635 | -82.1345
5.0638 | -82.1367
5.0641 | -82.1381
5.0644 | -82.1386
5.0648 | -82.14
5.0651 | -82.1414
5.0654 | -82.1429
5.0658 | -82.1454
5.0661 | -82.1474
5.0664 | -82.1482
5.0667 | -82.15
5.0671 | -82.1524
5.0674 | -82.1543
5.0677 | -82.1555
5.0681 | -82.1559
5.0684 | -82.1567
5.0687 | -82.1582
5.0691 | -82.1595
5.0694 | -82.1609
5.0697 | -82.1627
5.07 | -82.1642
5.0704 | -82.1653
5.0707 | -82.1663
5.071 | -82.1666
5.0714 | -82.1675
5.0717 | -82.1681
5.072 | -82.1676
5.0723 | -82.1678
5.0727 | -82.17
5.073 | -82.1724
5.0733 | -82.175
5.0737 | -82.1764
5.074 | -82.1767
5.0743 | -82.177
5.0746 | -82.1784
5.075 | -82.1798
5.0753 | -82.1818
5.0756 | -82.1836
5.076 | -82.1855
5.0763 | -82.1872
5.0766 | -82.1886
5.077 | -82.1901
5.0773 | -82.1914
5.0776 | -82.1925
5.0779 | -82.1933
5.0783 | -82.194
5.0786 | -82.195
5.0789 | -82.1963
5.0793 | -82.1978
5.0796 | -82.1992
5.0799 | -82.2005
5.0802 | -82.2019
5.0806 | -82.204
5.0809 | -82.2058
5.0812 | -82.2064
5.0816 | -82.2068
5.0819 | -82.2083
5.0822 | -82.2096
5.0825 | -82.2115
5.0829 | -82.2126
5.0832 | -82.2147
5.0835 | -82.2171
5.0839 | -82.2184
5.0842 | -82.2199
5.0845 | -82.2214
5.0849 | -82.2228
5.0852 | -82.2228
5.0855 | -82.2236
5.0858 | -82.225
5.0862 | -82.2266
5.0865 | -82.2279
5.0868 | -82.2293
5.0872 | -82.2308
5.0875 | -82.2329
5.0878 | -82.2344
5.0881 | -82.2356
5.0885 | -82.2368
5.0888 | -82.2387
5.0891 | -82.2405
5.0895 | -82.2413
5.0898 | -82.2427
5.0901 | -82.2444
5.0904 | -82.2465
5.0908 | -82.2488
5.0911 | -82.2518
5.0914 | -82.2541
5.0918 | -82.2551
5.0921 | -82.2567
5.0924 | -82.2582
5.0928 | -82.2599
5.0931 | -82.2617
5.0934 | -82.2628
5.0937 | -82.2637
5.0941 | -82.264
5.0944 | -82.2657
5.0947 | -82.2678
5.0951 | -82.2697
5.0954 | -82.2713
5.0957 | -82.2727
5.096 | -82.2745
5.0964 | -82.2765
5.0967 | -82.2793
5.097 | -82.281
5.0974 | -82.2825
5.0977 | -82.2829
5.098 | -82.2841
5.0983 | -82.2869
5.0987 | -82.2877
5.099 | -82.2895
5.0993 | -82.2918
5.0997 | -82.2935
5.1 | -82.2955
5.1003 | -82.2975
5.1006 | -82.2995
5.101 | -82.3006
5.1013 | -82.3012
5.1016 | -82.3028
5.102 | -82.3043
5.1023 | -82.3055
5.1026 | -82.3061
5.103 | -82.3075
5.1033 | -82.31
5.1036 | -82.312
5.1039 | -82.3147
5.1043 | -82.3164
5.1046 | -82.3178
5.1049 | -82.3199
5.1053 | -82.3217
5.1056 | -82.3233
5.1059 | -82.3247
5.1062 | -82.3262
5.1066 | -82.3269
5.1069 | -82.3281
5.1072 | -82.3295
5.1076 | -82.3303
5.1079 | -82.3313
5.1082 | -82.3323
5.1086 | -82.3342
5.1089 | -82.3362
5.1092 | -82.3379
5.1095 | -82.3395
5.1099 | -82.3416
5.1102 | -82.3446
5.1105 | -82.3474
5.1109 | -82.3499
5.1112 | -82.3514
5.1115 | -82.3542
5.1118 | -82.3572
5.1122 | -82.359
5.1125 | -82.3604
5.1128 | -82.3616
5.1132 | -82.3634
5.1135 | -82.3654
5.1138 | -82.367
5.1141 | -82.3686
5.1145 | -82.37
5.1148 | -82.3713
5.1151 | -82.3723
5.1155 | -82.3743
5.1158 | -82.3767
5.1161 | -82.3792
5.1164 | -82.3841
5.1168 | -82.3898
5.1171 | -82.3917
5.1174 | -82.3926
5.1178 | -82.3929
5.1181 | -82.3937
5.1184 | -82.3937
5.1188 | -82.3949
5.1191 | -82.3965
5.1194 | -82.3978
5.1197 | -82.3991
5.1201 | -82.4002
5.1204 | -82.4021
5.1207 | -82.4038
5.1211 | -82.4051
5.1214 | -82.4062
5.1217 | -82.4077
5.122 | -82.4092
5.1224 | -82.4114
5.1227 | -82.4133
5.123 | -82.4148
5.1234 | -82.4173
5.1237 | -82.4186
5.124 | -82.4207
5.1243 | -82.4225
5.1247 | -82.4243
5.125 | -82.4267
5.1253 | -82.4283
5.1257 | -82.4295
5.126 | -82.4308
5.1263 | -82.4324
5.1267 | -82.4346
5.127 | -82.4368
5.1273 | -82.439
5.1276 | -82.441
5.128 | -82.4422
5.1283 | -82.4436
5.1286 | -82.4459
5.129 | -82.4499
5.1293 | -82.4524
5.1296 | -82.4548
5.1299 | -82.4566
5.1303 | -82.4576
5.1306 | -82.4586
5.1309 | -82.4589
5.1313 | -82.459
5.1316 | -82.4603
5.1319 | -82.4615
5.1322 | -82.4632
5.1326 | -82.4644
5.1329 | -82.4659
5.1332 | -82.467
5.1336 | -82.468
5.1339 | -82.469
5.1342 | -82.4694
5.1346 | -82.4694
5.1349 | -82.4699
5.1352 | -82.471
5.1355 | -82.4717
5.1359 | -82.4721
5.1362 | -82.4722
5.1365 | -82.4738
5.1369 | -82.4753
5.1372 | -82.4761
5.1375 | -82.477
5.1378 | -82.4774
5.1382 | -82.4779
5.1385 | -82.4784
5.1388 | -82.4792
5.1392 | -82.4792
5.1395 | -82.4793
5.1398 | -82.48
5.1401 | -82.4807
5.1405 | -82.4807
5.1408 | -82.4811
5.1411 | -82.4821
5.1415 | -82.483
5.1418 | -82.4843
5.1421 | -82.4866
5.1425 | -82.4883
5.1428 | -82.4895
5.1431 | -82.4901
5.1434 | -82.491
5.1438 | -82.4923
5.1441 | -82.493
5.1444 | -82.4957
5.1448 | -82.4972
5.1451 | -82.4981
5.1454 | -82.4988
5.1457 | -82.4995
5.1461 | -82.5003
5.1464 | -82.5006
5.1467 | -82.5017
5.1471 | -82.5027
5.1474 | -82.5036
5.1477 | -82.5044
5.148 | -82.5044
5.1484 | -82.5049
5.1487 | -82.5055
5.149 | -82.5068
5.1494 | -82.5089
5.1497 | -82.5098
5.15 | -82.5102
5.1504 | -82.5107
5.1507 | -82.5116
5.151 | -82.5124
5.1513 | -82.5127
5.1517 | -82.5147
5.152 | -82.5162
5.1523 | -82.5183
5.1527 | -82.5215
5.153 | -82.5234
5.1533 | -82.5245
5.1536 | -82.5254
5.154 | -82.5266
5.1543 | -82.5276
5.1546 | -82.5277
5.155 | -82.5275
5.1553 | -82.5274
5.1556 | -82.5277
5.1559 | -82.5276
5.1563 | -82.5276
5.1566 | -82.528
5.1569 | -82.5291
5.1573 | -82.5298
5.1576 | -82.5311
5.1579 | -82.531
5.1582 | -82.5314
5.1586 | -82.5319
5.1589 | -82.5323
5.1592 | -82.5331
5.1596 | -82.5341
5.1599 | -82.5333
5.1602 | -82.5324
5.1606 | -82.5332
5.1609 | -82.534
5.1612 | -82.5344
5.1615 | -82.5341
5.1619 | -82.5341
5.1622 | -82.5343
5.1625 | -82.535
5.1629 | -82.5357
5.1632 | -82.5365
5.1635 | -82.5369
5.1638 | -82.5376
5.1642 | -82.5383
5.1645 | -82.5389
5.1648 | -82.5386
5.1652 | -82.5381
5.1655 | -82.5383
5.1658 | -82.5387
5.1661 | -82.5406
5.1665 | -82.5423
5.1668 | -82.5427
5.1671 | -82.543
5.1675 | -82.5432
5.1678 | -82.5436
5.1681 | -82.5433
5.1685 | -82.543
5.1688 | -82.5433
5.1691 | -82.5425
5.1694 | -82.5411
5.1698 | -82.541
5.1701 | -82.5418
5.1704 | -82.5423
5.1708 | -82.5427
5.1711 | -82.5436
5.1714 | -82.5439
5.1717 | -82.5443
5.1721 | -82.5442
5.1724 | -82.5435
5.1727 | -82.544
5.1731 | -82.5437
5.1734 | -82.5432
5.1737 | -82.5432
5.174 | -82.5435
5.1744 | -82.5435
5.1747 | -82.5442
5.175 | -82.5451
5.1754 | -82.5449
5.1757 | -82.5452
5.176 | -82.5461
5.1764 | -82.546
5.1767 | -82.5459
5.177 | -82.5464
5.1773 | -82.5461
5.1777 | -82.5439
5.178 | -82.5434
5.1783 | -82.5424
5.1787 | -82.541
5.179 | -82.5416
5.1793 | -82.5423
5.1796 | -82.5422
5.18 | -82.5423
5.1803 | -82.5429
5.1806 | -82.5438
5.181 | -82.5442
5.1813 | -82.5447
5.1816 | -82.5452
5.1819 | -82.5457
5.1823 | -82.5467
5.1826 | -82.5492
5.1829 | -82.5502
5.1833 | -82.5502
5.1836 | -82.5509
5.1839 | -82.5529
5.1843 | -82.5554
5.1846 | -82.5567
5.1849 | -82.5569
5.1852 | -82.5569
5.1856 | -82.5569
5.1859 | -82.5569
5.1862 | -82.5573
5.1866 | -82.5574
5.1869 | -82.5578
5.1872 | -82.5581
5.1875 | -82.5576
5.1879 | -82.5574
5.1882 | -82.5577
5.1885 | -82.5578
5.1889 | -82.5583
5.1892 | -82.5593
5.1895 | -82.5593
5.1898 | -82.5588
5.1902 | -82.5581
5.1905 | -82.558
5.1908 | -82.5569
5.1912 | -82.5564
5.1915 | -82.5564
5.1918 | -82.5562
5.1922 | -82.5562
5.1925 | -82.5562
5.1928 | -82.556
5.1931 | -82.555
5.1935 | -82.5545
5.1938 | -82.5547
5.1941 | -82.5554
5.1945 | -82.5552
5.1948 | -82.5554
5.1951 | -82.5553
5.1954 | -82.5552
5.1958 | -82.5551
5.1961 | -82.5554
5.1964 | -82.5562
5.1968 | -82.5565
5.1971 | -82.5566
5.1974 | -82.5571
5.1977 | -82.5571
5.1981 | -82.5572
5.1984 | -82.5578
5.1987 | -82.5572
5.1991 | -82.5572
5.1994 | -82.5579
5.1997 | -82.5581
5.2001 | -82.5572
5.2004 | -82.5563
5.2007 | -82.5564
5.201 | -82.5562
5.2014 | -82.5562
5.2017 | -82.5562
5.202 | -82.5555
5.2024 | -82.5556
5.2027 | -82.5574
5.203 | -82.5577
5.2033 | -82.5584
5.2037 | -82.5587
5.204 | -82.5583
5.2043 | -82.5584
5.2047 | -82.5583
5.205 | -82.5576
5.2053 | -82.5574
5.2056 | -82.5574
5.206 | -82.5573
5.2063 | -82.5573
5.2066 | -82.5569
5.207 | -82.5563
5.2073 | -82.5563
5.2076 | -82.556
5.2079 | -82.5567
5.2083 | -82.5573
5.2086 | -82.5572
5.2089 | -82.5569
5.2093 | -82.5563
5.2096 | -82.5559
5.2099 | -82.5558
5.2103 | -82.5564
5.2106 | -82.557
5.2109 | -82.5571
5.2112 | -82.5582
5.2116 | -82.559
5.2119 | -82.5596
5.2122 | -82.5593
5.2126 | -82.5582
5.2129 | -82.5574
5.2132 | -82.5565
5.2135 | -82.5562
5.2139 | -82.5554
5.2142 | -82.5554
5.2145 | -82.556
5.2149 | -82.5549
5.2152 | -82.5542
5.2155 | -82.5533
5.2158 | -82.5526
5.2162 | -82.5528
5.2165 | -82.5523
5.2168 | -82.5518
5.2172 | -82.5519
5.2175 | -82.5516
5.2178 | -82.5515
5.2182 | -82.5509
5.2185 | -82.5512
5.2188 | -82.5511
5.2191 | -82.5507
5.2195 | -82.5509
5.2198 | -82.5508
5.2201 | -82.5498
5.2205 | -82.5492
5.2208 | -82.5489
5.2211 | -82.5485
5.2214 | -82.5473
5.2218 | -82.5468
5.2221 | -82.5466
5.2224 | -82.5462
5.2228 | -82.5446
5.2231 | -82.5448
5.2234 | -82.5452
5.2237 | -82.5455
5.2241 | -82.5446
5.2244 | -82.5423
5.2247 | -82.541
5.2251 | -82.5421
5.2254 | -82.5431
5.2257 | -82.5429
5.2261 | -82.5413
5.2264 | -82.5407
5.2267 | -82.5405
5.227 | -82.5398
5.2274 | -82.5397
5.2277 | -82.5395
5.228 | -82.5393
5.2284 | -82.541
5.2287 | -82.5419
5.229 | -82.5413
5.2293 | -82.5413
5.2297 | -82.5415
5.23 | -82.5409
5.2303 | -82.5411
5.2307 | -82.5413
5.231 | -82.5409
5.2313 | -82.54
5.2316 | -82.5396
5.232 | -82.5394
5.2323 | -82.5392
5.2326 | -82.5389
5.233 | -82.5387
5.2333 | -82.5367
5.2336 | -82.5347
5.234 | -82.5345
5.2343 | -82.5338
5.2346 | -82.5336
5.2349 | -82.5342
5.2353 | -82.5347
5.2356 | -82.5347
5.2359 | -82.5344
5.2363 | -82.5338
5.2366 | -82.5332
5.2369 | -82.5328
5.2372 | -82.5326
5.2376 | -82.5323
5.2379 | -82.5333
5.2382 | -82.5337
5.2386 | -82.5333
5.2389 | -82.5329
5.2392 | -82.535
5.2395 | -82.5362
5.2399 | -82.5351
5.2402 | -82.5347
5.2405 | -82.534
5.2409 | -82.5324
5.2412 | -82.5288
5.2415 | -82.5254
5.2419 | -82.524
5.2422 | -82.5233
5.2425 | -82.523
5.2428 | -82.5215
5.2432 | -82.5197
5.2435 | -82.5185
5.2438 | -82.5175
5.2442 | -82.5171
5.2445 | -82.5171
5.2448 | -82.5174
5.2451 | -82.5166
5.2455 | -82.5154
5.2458 | -82.5147
5.2461 | -82.5139
5.2465 | -82.5131
5.2468 | -82.5126
5.2471 | -82.513
5.2474 | -82.5127
5.2478 | -82.5113
5.2481 | -82.5102
5.2484 | -82.5089
5.2488 | -82.5079
5.2491 | -82.5076
5.2494 | -82.5071
5.2497 | -82.5064
5.2501 | -82.5052
5.2504 | -82.5042
5.2507 | -82.5045
5.2511 | -82.5049
5.2514 | -82.505
5.2517 | -82.5055
5.2521 | -82.5059
5.2524 | -82.5056
5.2527 | -82.505
5.253 | -82.5042
5.2534 | -82.5034
5.2537 | -82.5026
5.254 | -82.5023
5.2544 | -82.5031
5.2547 | -82.5022
5.255 | -82.5019
5.2553 | -82.5018
5.2557 | -82.5016
5.256 | -82.5012
5.2563 | -82.5006
5.2567 | -82.5
5.257 | -82.4994
5.2573 | -82.4986
5.2576 | -82.4984
5.258 | -82.4989
5.2583 | -82.4994
5.2586 | -82.4986
5.259 | -82.4979
5.2593 | -82.4972
5.2596 | -82.4982
5.26 | -82.4999
5.2603 | -82.4995
5.2606 | -82.4978
5.2609 | -82.4974
5.2613 | -82.4965
5.2616 | -82.4949
5.2619 | -82.4946
5.2623 | -82.4947
5.2626 | -82.4952
5.2629 | -82.4944
5.2632 | -82.4927
5.2636 | -82.4904
5.2639 | -82.4884
5.2642 | -82.4886
5.2646 | -82.489
5.2649 | -82.4884
5.2652 | -82.4878
5.2655 | -82.4867
5.2659 | -82.4857
5.2662 | -82.4847
5.2665 | -82.484
5.2669 | -82.4835
5.2672 | -82.4829
5.2675 | -82.4824
5.2679 | -82.4817
5.2682 | -82.4812
5.2685 | -82.4798
5.2688 | -82.4789
5.2692 | -82.4783
5.2695 | -82.4777
5.2698 | -82.477
5.2702 | -82.4764
5.2705 | -82.4765
5.2708 | -82.476
5.2711 | -82.475
5.2715 | -82.4744
5.2718 | -82.4741
5.2721 | -82.4747
5.2725 | -82.4741
5.2728 | -82.4733
5.2731 | -82.4728
5.2734 | -82.4722
5.2738 | -82.4718
5.2741 | -82.4709
5.2744 | -82.4697
5.2748 | -82.4699
5.2751 | -82.4704
5.2754 | -82.4704
5.2758 | -82.4696
5.2761 | -82.4694
5.2764 | -82.4706
5.2767 | -82.4708
5.2771 | -82.4706
5.2774 | -82.47
5.2777 | -82.4695
5.2781 | -82.4695
5.2784 | -82.469
5.2787 | -82.4682
5.279 | -82.4681
5.2794 | -82.4684
5.2797 | -82.4692
5.28 | -82.4695
5.2804 | -82.4686
5.2807 | -82.4675
5.281 | -82.4673
5.2813 | -82.4679
5.2817 | -82.4676
5.282 | -82.467
5.2823 | -82.467
5.2827 | -82.4679
5.283 | -82.4674
5.2833 | -82.467
5.2837 | -82.4662
5.284 | -82.466
5.2843 | -82.4659
5.2846 | -82.4657
5.285 | -82.4659
5.2853 | -82.4655
5.2856 | -82.465
5.286 | -82.4633
5.2863 | -82.4619
5.2866 | -82.4613
5.2869 | -82.4609
5.2873 | -82.4605
5.2876 | -82.4602
5.2879 | -82.4598
5.2883 | -82.4597
5.2886 | -82.46
5.2889 | -82.4599
5.2892 | -82.4583
5.2896 | -82.4578
5.2899 | -82.4566
5.2902 | -82.4559
5.2906 | -82.4554
5.2909 | -82.4547
5.2912 | -82.4542
5.2915 | -82.4541
5.2919 | -82.4535
5.2922 | -82.4531
5.2925 | -82.4536
5.2929 | -82.4537
5.2932 | -82.4535
5.2935 | -82.4534
5.2939 | -82.4532
5.2942 | -82.4531
5.2945 | -82.4534
5.2948 | -82.4535
5.2952 | -82.4532
5.2955 | -82.4531
5.2958 | -82.453
5.2962 | -82.4532
5.2965 | -82.4539
5.2968 | -82.4555
5.2971 | -82.4559
5.2975 | -82.4557
5.2978 | -82.456
5.2981 | -82.456
5.2985 | -82.4559
5.2988 | -82.4559
5.2991 | -82.456
5.2995 | -82.456
5.2998 | -82.4562
5.3001 | -82.4564
5.3004 | -82.4559
5.3008 | -82.4554
5.3011 | -82.4559
5.3014 | -82.457
5.3018 | -82.4574
5.3021 | -82.4566
5.3024 | -82.4566
5.3027 | -82.4562
5.3031 | -82.4559
5.3034 | -82.4559
5.3037 | -82.4561
5.3041 | -82.456
5.3044 | -82.4562
5.3047 | -82.4567
5.305 | -82.4566
5.3054 | -82.4559
5.3057 | -82.4549
5.306 | -82.4546
5.3064 | -82.4553
5.3067 | -82.4552
5.307 | -82.4561
5.3073 | -82.4576
5.3077 | -82.4576
5.308 | -82.4585
5.3083 | -82.4593
5.3087 | -82.4596
5.309 | -82.4598
5.3093 | -82.4594
5.3097 | -82.4594
5.31 | -82.4597
5.3103 | -82.4604
5.3106 | -82.4604
5.311 | -82.4605
5.3113 | -82.4611
5.3116 | -82.4611
5.312 | -82.4618
5.3123 | -82.4622
5.3126 | -82.4624
5.3129 | -82.4622
5.3133 | -82.4619
5.3136 | -82.4619
5.3139 | -82.4624
5.3143 | -82.4632
5.3146 | -82.4635
5.3149 | -82.4638
5.3152 | -82.4644
5.3156 | -82.4644
5.3159 | -82.4643
5.3162 | -82.4636
5.3166 | -82.464
5.3169 | -82.4638
5.3172 | -82.4628
5.3176 | -82.4622
5.3179 | -82.4629
5.3182 | -82.4623
5.3185 | -82.4623
5.3189 | -82.4629
5.3192 | -82.4638
5.3195 | -82.4642
5.3199 | -82.4646
5.3202 | -82.4649
5.3205 | -82.4648
5.3208 | -82.4652
5.3212 | -82.4653
5.3215 | -82.4652
5.3218 | -82.4659
5.3222 | -82.4663
5.3225 | -82.4666
5.3228 | -82.4669
5.3231 | -82.4679
5.3235 | -82.4688
5.3238 | -82.4689
5.3241 | -82.4694
5.3245 | -82.4693
5.3248 | -82.4687
5.3251 | -82.4692
5.3255 | -82.4701
5.3258 | -82.4705
5.3261 | -82.4707
5.3264 | -82.4712
5.3268 | -82.472
5.3271 | -82.4724
5.3274 | -82.4739
5.3278 | -82.4742
5.3281 | -82.4755
5.3284 | -82.4767
5.3287 | -82.4772
5.3291 | -82.4781
5.3294 | -82.4787
5.3297 | -82.4792
5.3301 | -82.4795
5.3304 | -82.479
5.3307 | -82.4778
5.331 | -82.4779
5.3314 | -82.4788
5.3317 | -82.4796
5.332 | -82.4804
5.3324 | -82.4813
5.3327 | -82.4818
5.333 | -82.4825
5.3334 | -82.4825
5.3337 | -82.4831
5.334 | -82.4836
5.3343 | -82.4846
5.3347 | -82.4857
5.335 | -82.4868
5.3353 | -82.4876
5.3357 | -82.4886
5.336 | -82.4897
5.3363 | -82.4904
5.3366 | -82.4903
5.337 | -82.4919
5.3373 | -82.4922
5.3376 | -82.4929
5.338 | -82.4935
5.3383 | -82.4942
5.3386 | -82.4939
5.3389 | -82.4943
5.3393 | -82.495
5.3396 | -82.4956
5.3399 | -82.497
5.3403 | -82.4986
5.3406 | -82.4988
5.3409 | -82.4964
5.3412 | -82.495
5.3416 | -82.4973
5.3419 | -82.5001
5.3422 | -82.5007
5.3426 | -82.5014
5.3429 | -82.5018
5.3432 | -82.5017
5.3436 | -82.5019
5.3439 | -82.5023
5.3442 | -82.504
5.3445 | -82.5054
5.3449 | -82.5065
5.3452 | -82.5072
5.3455 | -82.5079
5.3459 | -82.5088
5.3462 | -82.51
5.3465 | -82.5098
5.3468 | -82.5097
5.3472 | -82.5105
5.3475 | -82.5113
5.3478 | -82.5113
5.3482 | -82.5111
5.3485 | -82.5115
5.3488 | -82.5122
5.3491 | -82.5131
5.3495 | -82.5134
5.3498 | -82.514
5.3501 | -82.5149
5.3505 | -82.5161
5.3508 | -82.5172
5.3511 | -82.5181
5.3515 | -82.5196
5.3518 | -82.5209
5.3521 | -82.5209
5.3524 | -82.5216
5.3528 | -82.5229
5.3531 | -82.5234
5.3534 | -82.5239
5.3538 | -82.525
5.3541 | -82.5264
5.3544 | -82.527
5.3547 | -82.5273
5.3551 | -82.5279
5.3554 | -82.5287
5.3557 | -82.5301
5.3561 | -82.5298
5.3564 | -82.5304
5.3567 | -82.5306
5.357 | -82.5312
5.3574 | -82.5315
5.3577 | -82.5318
5.358 | -82.5344
5.3584 | -82.5359
5.3587 | -82.5376
5.359 | -82.5394
5.3594 | -82.5406
5.3597 | -82.5416
5.36 | -82.5421
5.3603 | -82.5433
5.3607 | -82.5448
5.361 | -82.5452
5.3613 | -82.5455
5.3617 | -82.5465
5.362 | -82.5471
5.3623 | -82.5467
5.3626 | -82.5469
5.363 | -82.5483
5.3633 | -82.5486
5.3636 | -82.5487
5.364 | -82.5497
5.3643 | -82.5512
5.3646 | -82.5518
5.3649 | -82.5516
5.3653 | -82.5536
5.3656 | -82.5545
5.3659 | -82.5553
5.3663 | -82.5566
5.3666 | -82.5576
5.3669 | -82.5588
5.3673 | -82.56
5.3676 | -82.5602
5.3679 | -82.5602
5.3682 | -82.5613
5.3686 | -82.5625
5.3689 | -82.5637
5.3692 | -82.5645
5.3696 | -82.5651
5.3699 | -82.565
5.3702 | -82.5655
5.3705 | -82.5666
5.3709 | -82.566
5.3712 | -82.5661
5.3715 | -82.5665
5.3719 | -82.5672
5.3722 | -82.5691
5.3725 | -82.5702
5.3728 | -82.5711
5.3732 | -82.572
5.3735 | -82.5729
5.3738 | -82.5738
5.3742 | -82.5749
5.3745 | -82.5762
5.3748 | -82.5772
5.3752 | -82.5784
5.3755 | -82.5793
5.3758 | -82.5798
5.3761 | -82.5809
5.3765 | -82.5828
5.3768 | -82.5832
5.3771 | -82.5837
5.3775 | -82.5847
5.3778 | -82.5856
5.3781 | -82.5866
5.3784 | -82.5869
5.3788 | -82.5871
5.3791 | -82.5884
5.3794 | -82.5895
5.3798 | -82.5903
5.3801 | -82.5904
5.3804 | -82.5908
5.3807 | -82.5917
5.3811 | -82.5926
5.3814 | -82.5935
5.3817 | -82.5941
5.3821 | -82.5948
5.3824 | -82.5958
5.3827 | -82.597
5.3831 | -82.5983
5.3834 | -82.5996
5.3837 | -82.6013
5.384 | -82.6026
5.3844 | -82.604
5.3847 | -82.6048
5.385 | -82.6064
5.3854 | -82.6076
5.3857 | -82.6073
5.386 | -82.6079
5.3863 | -82.6086
5.3867 | -82.6092
5.387 | -82.6101
5.3873 | -82.6111
5.3877 | -82.612
5.388 | -82.6129
5.3883 | -82.6135
5.3886 | -82.6122
5.389 | -82.6127
5.3893 | -82.6134
5.3896 | -82.6143
5.39 | -82.6155
5.3903 | -82.6161
5.3906 | -82.6171
5.391 | -82.618
5.3913 | -82.6182
5.3916 | -82.6191
5.3919 | -82.6213
5.3923 | -82.6233
5.3926 | -82.625
5.3929 | -82.6256
5.3933 | -82.6267
5.3936 | -82.6293
5.3939 | -82.6306
5.3942 | -82.6318
5.3946 | -82.633
5.3949 | -82.6338
5.3952 | -82.6346
5.3956 | -82.6356
5.3959 | -82.6367
5.3962 | -82.6365
5.3965 | -82.6357
5.3969 | -82.6364
5.3972 | -82.6371
5.3975 | -82.6368
5.3979 | -82.6365
5.3982 | -82.6369
5.3985 | -82.6379
5.3988 | -82.6389
5.3992 | -82.6402
5.3995 | -82.6414
5.3998 | -82.6428
5.4002 | -82.644
5.4005 | -82.6441
5.4008 | -82.6437
5.4012 | -82.6441
5.4015 | -82.6461
5.4018 | -82.6479
5.4021 | -82.649
5.4025 | -82.651
5.4028 | -82.6531
5.4031 | -82.6543
5.4035 | -82.6559
5.4038 | -82.6575
5.4041 | -82.662
5.4044 | -82.6634
5.4048 | -82.6632
5.4051 | -82.6636
5.4054 | -82.6639
5.4058 | -82.6641
5.4061 | -82.6648
5.4064 | -82.6657
5.4067 | -82.6668
5.4071 | -82.6686
5.4074 | -82.6698
5.4077 | -82.6702
5.4081 | -82.6707
5.4084 | -82.6714
5.4087 | -82.6723
5.4091 | -82.6731
5.4094 | -82.6746
5.4097 | -82.6755
5.41 | -82.6768
5.4104 | -82.6779
5.4107 | -82.6783
5.411 | -82.6782
5.4114 | -82.679
5.4117 | -82.6805
5.412 | -82.6823
5.4123 | -82.6828
5.4127 | -82.6813
5.413 | -82.6825
5.4133 | -82.6839
5.4137 | -82.6859
5.414 | -82.6875
5.4143 | -82.6884
5.4146 | -82.6893
5.415 | -82.6901
5.4153 | -82.6911
5.4156 | -82.6918
5.416 | -82.6919
5.4163 | -82.6926
5.4166 | -82.6935
5.417 | -82.6939
5.4173 | -82.6938
5.4176 | -82.6938
5.4179 | -82.6942
5.4183 | -82.6951
5.4186 | -82.6971
5.4189 | -82.6981
5.4193 | -82.6985
5.4196 | -82.6988
5.4199 | -82.6985
5.4202 | -82.7001
5.4206 | -82.7023
5.4209 | -82.7037
5.4212 | -82.7053
5.4216 | -82.7063
5.4219 | -82.7069
5.4222 | -82.7056
5.4225 | -82.7057
5.4229 | -82.7064
5.4232 | -82.707
5.4235 | -82.7072
5.4239 | -82.7076
5.4242 | -82.7086
5.4245 | -82.7104
5.4249 | -82.712
5.4252 | -82.7123
5.4255 | -82.7126
5.4258 | -82.7136
5.4262 | -82.7145
5.4265 | -82.7153
5.4268 | -82.7165
5.4272 | -82.7175
5.4275 | -82.7175
5.4278 | -82.718
5.4281 | -82.7186
5.4285 | -82.72
5.4288 | -82.7224
5.4291 | -82.724
5.4295 | -82.7246
5.4298 | -82.7253
5.4301 | -82.7257
5.4304 | -82.7265
5.4308 | -82.7275
5.4311 | -82.7278
5.4314 | -82.7284
5.4318 | -82.7291
5.4321 | -82.7296
5.4324 | -82.7303
5.4328 | -82.732
5.4331 | -82.7335
5.4334 | -82.7339
5.4337 | -82.7343
5.4341 | -82.7346
5.4344 | -82.7355
5.4347 | -82.736
5.4351 | -82.7367
5.4354 | -82.7381
5.4357 | -82.7395
5.436 | -82.7404
5.4364 | -82.7409
5.4367 | -82.7424
5.437 | -82.7451
5.4374 | -82.7474
5.4377 | -82.7479
5.438 | -82.7486
5.4383 | -82.7489
5.4387 | -82.7496
5.439 | -82.7512
5.4393 | -82.7534
5.4397 | -82.7552
5.44 | -82.7562
5.4403 | -82.7568
5.4406 | -82.7574
5.441 | -82.7584
5.4413 | -82.7584
5.4416 | -82.7597
5.442 | -82.7608
5.4423 | -82.761
5.4426 | -82.7616
5.443 | -82.7626
5.4433 | -82.7636
5.4436 | -82.764
5.4439 | -82.764
5.4443 | -82.7642
5.4446 | -82.7646
5.4449 | -82.7653
5.4453 | -82.7658
5.4456 | -82.766
5.4459 | -82.7665
5.4462 | -82.7668
5.4466 | -82.7662
5.4469 | -82.7672
5.4472 | -82.769
5.4476 | -82.7695
5.4479 | -82.77
5.4482 | -82.7706
5.4485 | -82.7715
5.4489 | -82.7721
5.4492 | -82.7722
5.4495 | -82.7729
5.4499 | -82.7736
5.4502 | -82.7736
5.4505 | -82.7744
5.4509 | -82.7746
5.4512 | -82.7764
5.4515 | -82.7767
5.4518 | -82.776
5.4522 | -82.7758
5.4525 | -82.7756
5.4528 | -82.7747
5.4532 | -82.774
5.4535 | -82.7742
5.4538 | -82.7752
5.4541 | -82.7758
5.4545 | -82.7761
5.4548 | -82.7763
5.4551 | -82.777
5.4555 | -82.7785
5.4558 | -82.7795
5.4561 | -82.7799
5.4564 | -82.7798
5.4568 | -82.7801
5.4571 | -82.7803
5.4574 | -82.7799
5.4578 | -82.7801
5.4581 | -82.781
5.4584 | -82.7813
5.4588 | -82.7818
5.4591 | -82.7815
5.4594 | -82.7818
5.4597 | -82.7823
5.4601 | -82.7834
5.4604 | -82.7835
5.4607 | -82.7839
5.4611 | -82.7838
5.4614 | -82.7836
5.4617 | -82.7839
5.462 | -82.7841
5.4624 | -82.7829
5.4627 | -82.7814
5.463 | -82.7812
5.4634 | -82.7814
5.4637 | -82.7817
5.464 | -82.7818
5.4643 | -82.7832
5.4647 | -82.7839
5.465 | -82.7843
5.4653 | -82.7847
5.4657 | -82.7848
5.466 | -82.7847
5.4663 | -82.7848
5.4667 | -82.7849
5.467 | -82.7852
5.4673 | -82.7855
5.4676 | -82.7854
5.468 | -82.7859
5.4683 | -82.7868
5.4686 | -82.7867
5.469 | -82.787
5.4693 | -82.7877
5.4696 | -82.7888
5.4699 | -82.7882
5.4703 | -82.7884
5.4706 | -82.7888
5.4709 | -82.7892
5.4713 | -82.7888
5.4716 | -82.7882
5.4719 | -82.7886
5.4722 | -82.7873
5.4726 | -82.7861
5.4729 | -82.786
5.4732 | -82.7863
5.4736 | -82.7864
5.4739 | -82.7873
5.4742 | -82.7877
5.4746 | -82.7884
5.4749 | -82.7897
5.4752 | -82.7906
5.4755 | -82.7915
5.4759 | -82.7919
5.4762 | -82.7919
5.4765 | -82.7919
5.4769 | -82.7923
5.4772 | -82.7922
5.4775 | -82.7905
5.4778 | -82.7884
5.4782 | -82.788
5.4785 | -82.7876
5.4788 | -82.7879
5.4792 | -82.7878
5.4795 | -82.7868
5.4798 | -82.7864
5.4801 | -82.786
5.4805 | -82.7857
5.4808 | -82.7861
5.4811 | -82.7859
5.4815 | -82.7858
5.4818 | -82.786
5.4821 | -82.7859
5.4825 | -82.7856
5.4828 | -82.7853
5.4831 | -82.7859
5.4834 | -82.7862
5.4838 | -82.7858
5.4841 | -82.7857
5.4844 | -82.7858
5.4848 | -82.7854
5.4851 | -82.7854
5.4854 | -82.7858
5.4857 | -82.7867
5.4861 | -82.7875
5.4864 | -82.7877
5.4867 | -82.7882
5.4871 | -82.7888
5.4874 | -82.7884
5.4877 | -82.7877
5.488 | -82.787
5.4884 | -82.7872
5.4887 | -82.7872
5.489 | -82.787
5.4894 | -82.7871
5.4897 | -82.7873
5.49 | -82.7886
5.4904 | -82.7899
5.4907 | -82.7906
5.491 | -82.7912
5.4913 | -82.7915
5.4917 | -82.7914
5.492 | -82.7914
5.4923 | -82.7913
5.4927 | -82.7913
5.493 | -82.7916
5.4933 | -82.7917
5.4936 | -82.7908
5.494 | -82.7902
5.4943 | -82.79
5.4946 | -82.7885
5.495 | -82.7883
5.4953 | -82.7887
5.4956 | -82.7893
5.4959 | -82.7895
5.4963 | -82.7891
5.4966 | -82.7888
5.4969 | -82.7884
5.4973 | -82.788
5.4976 | -82.788
5.4979 | -82.7879
5.4982 | -82.7882
5.4986 | -82.7891
5.4989 | -82.7893
5.4992 | -82.7892
5.4996 | -82.7891
5.4999 | -82.7894
5.5002 | -82.7907
5.5006 | -82.7914
5.5009 | -82.7914
5.5012 | -82.7911
5.5015 | -82.7908
5.5019 | -82.7909
5.5022 | -82.7911
5.5025 | -82.7915
5.5029 | -82.7917
5.5032 | -82.792
5.5035 | -82.7927
5.5038 | -82.7938
5.5042 | -82.794
5.5045 | -82.7938
5.5048 | -82.7936
5.5052 | -82.7943
5.5055 | -82.7965
5.5058 | -82.7963
5.5061 | -82.7948
5.5065 | -82.7947
5.5068 | -82.7949
5.5071 | -82.7949
5.5075 | -82.7948
5.5078 | -82.794
5.5081 | -82.7953
5.5085 | -82.7962
5.5088 | -82.7961
5.5091 | -82.7964
5.5094 | -82.7971
5.5098 | -82.7971
5.5101 | -82.7973
5.5104 | -82.7979
5.5108 | -82.7983
5.5111 | -82.7981
5.5114 | -82.7975
5.5117 | -82.7968
5.5121 | -82.7966
5.5124 | -82.7966
5.5127 | -82.7966
5.5131 | -82.7958
5.5134 | -82.7946
5.5137 | -82.7944
5.514 | -82.7946
5.5144 | -82.7941
5.5147 | -82.7935
5.515 | -82.7932
5.5154 | -82.7932
5.5157 | -82.7935
5.516 | -82.7939
5.5164 | -82.7943
5.5167 | -82.7935
5.517 | -82.7928
5.5173 | -82.7933
5.5177 | -82.7934
5.518 | -82.7934
5.5183 | -82.7934
5.5187 | -82.7941
5.519 | -82.7946
5.5193 | -82.7948
5.5196 | -82.7949
5.52 | -82.7949
5.5203 | -82.7949
5.5206 | -82.795
5.521 | -82.7944
5.5213 | -82.7935
5.5216 | -82.7928
5.5219 | -82.7937
5.5223 | -82.794
5.5226 | -82.7938
5.5229 | -82.7936
5.5233 | -82.7936
5.5236 | -82.7936
5.5239 | -82.7935
5.5243 | -82.7936
5.5246 | -82.7937
5.5249 | -82.7937
5.5252 | -82.7942
5.5256 | -82.7952
5.5259 | -82.796
5.5262 | -82.7967
5.5266 | -82.7977
5.5269 | -82.7985
5.5272 | -82.7987
5.5275 | -82.7991
5.5279 | -82.799
5.5282 | -82.7992
5.5285 | -82.7991
5.5289 | -82.7992
5.5292 | -82.8
5.5295 | -82.8007
5.5298 | -82.8009
5.5302 | -82.8002
5.5305 | -82.7999
5.5308 | -82.8003
5.5312 | -82.8021
5.5315 | -82.804
5.5318 | -82.8042
5.5321 | -82.8035
5.5325 | -82.8036
5.5328 | -82.8047
5.5331 | -82.8051
5.5335 | -82.8053
5.5338 | -82.8051
5.5341 | -82.8046
5.5345 | -82.8038
5.5348 | -82.8038
5.5351 | -82.8038
5.5354 | -82.8037
5.5358 | -82.8035
5.5361 | -82.8043
5.5364 | -82.8051
5.5368 | -82.8054
5.5371 | -82.8045
5.5374 | -82.8022
5.5377 | -82.8011
5.5381 | -82.801
5.5384 | -82.8011
5.5387 | -82.8012
5.5391 | -82.8011
5.5394 | -82.8016
5.5397 | -82.8018
5.54 | -82.8025
5.5404 | -82.8029
5.5407 | -82.8028
5.541 | -82.8038
5.5414 | -82.8054
5.5417 | -82.8055
5.542 | -82.8052
5.5424 | -82.8051
5.5427 | -82.8053
5.543 | -82.8052
5.5433 | -82.8051
5.5437 | -82.8057
5.544 | -82.8063
5.5443 | -82.8071
5.5447 | -82.8067
5.545 | -82.8077
5.5453 | -82.8086
5.5456 | -82.8076
5.546 | -82.8077
5.5463 | -82.8076
5.5466 | -82.8081
5.547 | -82.8084
5.5473 | -82.809
5.5476 | -82.809
5.5479 | -82.8074
5.5483 | -82.8061
5.5486 | -82.8062
5.5489 | -82.8067
5.5493 | -82.8066
5.5496 | -82.8062
5.5499 | -82.8067
5.5503 | -82.8086
5.5506 | -82.8096
5.5509 | -82.8106
5.5512 | -82.8113
5.5516 | -82.8104
5.5519 | -82.8091
5.5522 | -82.8086
5.5526 | -82.8087
5.5529 | -82.8086
5.5532 | -82.8085
5.5535 | -82.808
5.5539 | -82.8076
5.5542 | -82.8074
5.5545 | -82.8074
5.5549 | -82.8078
5.5552 | -82.8089
5.5555 | -82.809
5.5558 | -82.809
5.5562 | -82.8091
5.5565 | -82.8092
5.5568 | -82.8093
5.5572 | -82.8089
5.5575 | -82.8087
5.5578 | -82.8089
5.5582 | -82.8088
5.5585 | -82.8089
5.5588 | -82.809
5.5591 | -82.809
5.5595 | -82.8089
5.5598 | -82.8088
5.5601 | -82.8088
5.5605 | -82.8089
5.5608 | -82.8095
5.5611 | -82.8102
5.5614 | -82.8108
5.5618 | -82.8114
5.5621 | -82.8117
5.5624 | -82.8116
5.5628 | -82.8115
5.5631 | -82.8117
5.5634 | -82.8125
5.5637 | -82.8131
5.5641 | -82.8144
5.5644 | -82.8151
5.5647 | -82.8165
5.5651 | -82.817
5.5654 | -82.8173
5.5657 | -82.8172
5.5661 | -82.8168
5.5664 | -82.8175
5.5667 | -82.8191
5.567 | -82.8197
5.5674 | -82.8199
5.5677 | -82.8202
5.568 | -82.8208
5.5684 | -82.8203
5.5687 | -82.8201
5.569 | -82.8204
5.5693 | -82.82
5.5697 | -82.8186
5.57 | -82.8174
5.5703 | -82.8174
5.5707 | -82.8183
5.571 | -82.8189
5.5713 | -82.8186
5.5716 | -82.8183
5.572 | -82.8183
5.5723 | -82.8187
5.5726 | -82.8194
5.573 | -82.8197
5.5733 | -82.8204
5.5736 | -82.8207
5.574 | -82.8218
5.5743 | -82.8219
5.5746 | -82.8233
5.5749 | -82.8236
5.5753 | -82.8238
5.5756 | -82.824
5.5759 | -82.8242
5.5763 | -82.8248
5.5766 | -82.8251
5.5769 | -82.8255
5.5772 | -82.826
5.5776 | -82.8263
5.5779 | -82.8269
5.5782 | -82.8278
5.5786 | -82.8278
5.5789 | -82.828
5.5792 | -82.8285
5.5795 | -82.8292
5.5799 | -82.8296
5.5802 | -82.8302
5.5805 | -82.8306
5.5809 | -82.8309
5.5812 | -82.8305
5.5815 | -82.8304
5.5819 | -82.8298
5.5822 | -82.8286
5.5825 | -82.8278
5.5828 | -82.8278
5.5832 | -82.8282
5.5835 | -82.8289
5.5838 | -82.8298
5.5842 | -82.8305
5.5845 | -82.8309
5.5848 | -82.8311
5.5851 | -82.8317
5.5855 | -82.8325
5.5858 | -82.8338
5.5861 | -82.8343
5.5865 | -82.8342
5.5868 | -82.8344
5.5871 | -82.8355
5.5874 | -82.8368
5.5878 | -82.8373
5.5881 | -82.8382
5.5884 | -82.8384
5.5888 | -82.839
5.5891 | -82.8396
5.5894 | -82.841
5.5897 | -82.842
5.5901 | -82.8427
5.5904 | -82.8436
5.5907 | -82.8442
5.5911 | -82.8443
5.5914 | -82.8442
5.5917 | -82.845
5.5921 | -82.8458
5.5924 | -82.8462
5.5927 | -82.8469
5.593 | -82.8476
5.5934 | -82.8489
5.5937 | -82.8496
5.594 | -82.8489
5.5944 | -82.848
5.5947 | -82.8486
5.595 | -82.8492
5.5953 | -82.8499
5.5957 | -82.8512
5.596 | -82.8514
5.5963 | -82.8515
5.5967 | -82.8529
5.597 | -82.8545
5.5973 | -82.8551
5.5976 | -82.8557
5.598 | -82.8564
5.5983 | -82.8581
5.5986 | -82.8591
5.599 | -82.8596
5.5993 | -82.8591
5.5996 | -82.8586
5.6 | -82.8601
5.6003 | -82.862
5.6006 | -82.8629
5.6009 | -82.8631
5.6013 | -82.8627
5.6016 | -82.8624
5.6019 | -82.8636
5.6023 | -82.8643
5.6026 | -82.8639
5.6029 | -82.8638
5.6032 | -82.8643
5.6036 | -82.8651
5.6039 | -82.8655
5.6042 | -82.8659
5.6046 | -82.8668
5.6049 | -82.8674
5.6052 | -82.8671
5.6055 | -82.8667
5.6059 | -82.8668
5.6062 | -82.867
5.6065 | -82.8667
5.6069 | -82.8669
5.6072 | -82.8674
5.6075 | -82.8679
5.6079 | -82.8683
5.6082 | -82.8688
5.6085 | -82.8693
5.6088 | -82.8702
5.6092 | -82.8712
5.6095 | -82.8724
5.6098 | -82.8746
5.6102 | -82.8761
5.6105 | -82.8766
5.6108 | -82.8776
5.6111 | -82.8776
5.6115 | -82.8775
5.6118 | -82.8781
5.6121 | -82.8793
5.6125 | -82.8795
5.6128 | -82.8798
5.6131 | -82.8803
5.6134 | -82.8809
5.6138 | -82.8813
5.6141 | -82.8818
5.6144 | -82.8826
5.6148 | -82.8834
5.6151 | -82.8841
5.6154 | -82.8847
5.6158 | -82.8853
5.6161 | -82.8853
5.6164 | -82.8852
5.6167 | -82.8867
5.6171 | -82.8888
5.6174 | -82.8902
5.6177 | -82.8918
5.6181 | -82.8953
5.6184 | -82.8986
5.6187 | -82.8999
5.619 | -82.901
5.6194 | -82.9017
5.6197 | -82.902
5.62 | -82.9028
5.6204 | -82.9041
5.6207 | -82.9052
5.621 | -82.9055
5.6213 | -82.9055
5.6217 | -82.9065
5.622 | -82.9069
5.6223 | -82.9068
5.6227 | -82.9075
5.623 | -82.9085
5.6233 | -82.909
5.6237 | -82.9096
5.624 | -82.9101
5.6243 | -82.9105
5.6246 | -82.9109
5.625 | -82.9115
5.6253 | -82.9124
5.6256 | -82.9135
5.626 | -82.9146
5.6263 | -82.9157
5.6266 | -82.9166
5.6269 | -82.9169
5.6273 | -82.9167
5.6276 | -82.9173
5.6279 | -82.9182
5.6283 | -82.9188
5.6286 | -82.9193
5.6289 | -82.9188
5.6292 | -82.9189
5.6296 | -82.9194
5.6299 | -82.9201
5.6302 | -82.9209
5.6306 | -82.9215
5.6309 | -82.922
5.6312 | -82.9214
5.6315 | -82.9226
5.6319 | -82.9232
5.6322 | -82.9243
5.6325 | -82.9247
5.6329 | -82.9255
5.6332 | -82.9266
5.6335 | -82.9268
5.6339 | -82.9272
5.6342 | -82.9278
5.6345 | -82.9279
5.6348 | -82.9281
5.6352 | -82.9283
5.6355 | -82.9286
5.6358 | -82.9291
5.6362 | -82.9297
5.6365 | -82.9304
5.6368 | -82.9308
5.6371 | -82.9309
5.6375 | -82.9318
5.6378 | -82.9325
5.6381 | -82.9328
5.6385 | -82.9334
5.6388 | -82.9351
5.6391 | -82.9362
5.6394 | -82.9359
5.6398 | -82.9368
5.6401 | -82.9383
5.6404 | -82.939
5.6408 | -82.9398
5.6411 | -82.9398
5.6414 | -82.9406
5.6418 | -82.9415
5.6421 | -82.9434
5.6424 | -82.9459
5.6427 | -82.948
5.6431 | -82.9492
5.6434 | -82.9495
5.6437 | -82.9498
5.6441 | -82.9504
5.6444 | -82.9515
5.6447 | -82.9525
5.645 | -82.9531
5.6454 | -82.9536
5.6457 | -82.9537
5.646 | -82.9538
5.6464 | -82.9533
5.6467 | -82.9538
5.647 | -82.9544
5.6473 | -82.9558
5.6477 | -82.9567
5.648 | -82.9573
5.6483 | -82.9582
5.6487 | -82.9591
5.649 | -82.96
5.6493 | -82.9609
5.6497 | -82.9614
5.65 | -82.9615
5.6503 | -82.9625
5.6506 | -82.9639
5.651 | -82.9649
5.6513 | -82.9653
5.6516 | -82.9656
5.652 | -82.9664
5.6523 | -82.9671
5.6526 | -82.9671
5.6529 | -82.9668
5.6533 | -82.9671
5.6536 | -82.9682
5.6539 | -82.9697
5.6543 | -82.9703
5.6546 | -82.971
5.6549 | -82.972
5.6552 | -82.973
5.6556 | -82.9737
5.6559 | -82.974
5.6562 | -82.9748
5.6566 | -82.976
5.6569 | -82.9769
5.6572 | -82.9768
5.6576 | -82.9771
5.6579 | -82.978
5.6582 | -82.979
5.6585 | -82.9798
5.6589 | -82.9802
5.6592 | -82.9799
5.6595 | -82.9798
5.6599 | -82.9797
5.6602 | -82.9803
5.6605 | -82.9816
5.6608 | -82.9832
5.6612 | -82.9844
5.6615 | -82.9853
5.6618 | -82.9865
5.6622 | -82.9868
5.6625 | -82.9867
5.6628 | -82.9876
5.6631 | -82.9883
5.6635 | -82.989
5.6638 | -82.9902
5.6641 | -82.9905
5.6645 | -82.9912
5.6648 | -82.9922
5.6651 | -82.9927
5.6655 | -82.9933
5.6658 | -82.9941
5.6661 | -82.9961
5.6664 | -82.9978
5.6668 | -82.9987
5.6671 | -82.9992
5.6674 | -82.9998
5.6678 | -83.0005
5.6681 | -83.0013
5.6684 | -83.0019
5.6687 | -83.0027
5.6691 | -83.003
5.6694 | -83.0036
5.6697 | -83.006
5.6701 | -83.0076
5.6704 | -83.0088
5.6707 | -83.0105
5.671 | -83.0116
5.6714 | -83.0126
5.6717 | -83.0136
5.672 | -83.0141
5.6724 | -83.0149
5.6727 | -83.0157
5.673 | -83.0163
5.6734 | -83.017
5.6737 | -83.0176
5.674 | -83.0177
5.6743 | -83.0181
5.6747 | -83.0185
5.675 | -83.0191
5.6753 | -83.0198
5.6757 | -83.0199
5.676 | -83.0199
5.6763 | -83.0188
5.6766 | -83.0186
5.677 | -83.0192
5.6773 | -83.0198
5.6776 | -83.0205
5.678 | -83.0214
5.6783 | -83.0213
5.6786 | -83.022
5.6789 | -83.0217
5.6793 | -83.022
5.6796 | -83.0218
5.6799 | -83.0223
5.6803 | -83.0231
5.6806 | -83.0242
5.6809 | -83.0249
5.6813 | -83.0251
5.6816 | -83.026
5.6819 | -83.0275
5.6822 | -83.0283
5.6826 | -83.0268
5.6829 | -83.0259
5.6832 | -83.0264
5.6836 | -83.0277
5.6839 | -83.0272
5.6842 | -83.0274
5.6845 | -83.0283
5.6849 | -83.0292
5.6852 | -83.0298
5.6855 | -83.0306
5.6859 | -83.0311
5.6862 | -83.0315
5.6865 | -83.0319
5.6868 | -83.0333
5.6872 | -83.0347
5.6875 | -83.0345
5.6878 | -83.0341
5.6882 | -83.0327
5.6885 | -83.0331
5.6888 | -83.0335
5.6891 | -83.034
5.6895 | -83.0343
5.6898 | -83.0346
5.6901 | -83.0353
5.6905 | -83.036
5.6908 | -83.0358
5.6911 | -83.0354
5.6915 | -83.0359
5.6918 | -83.0367
5.6921 | -83.0375
5.6924 | -83.0384
5.6928 | -83.0393
5.6931 | -83.0397
5.6934 | -83.0397
5.6938 | -83.04
5.6941 | -83.0403
5.6944 | -83.0406
5.6947 | -83.0411
5.6951 | -83.0419
5.6954 | -83.0425
5.6957 | -83.0427
5.6961 | -83.0428
5.6964 | -83.0429
5.6967 | -83.0431
5.697 | -83.0439
5.6974 | -83.0443
5.6977 | -83.0449
5.698 | -83.0455
5.6984 | -83.0467
5.6987 | -83.0482
5.699 | -83.0489
5.6994 | -83.05
5.6997 | -83.0519
5.7 | -83.0533
5.7003 | -83.0538
5.7007 | -83.0548
5.701 | -83.056
5.7013 | -83.0567
5.7017 | -83.0572
5.702 | -83.0576
5.7023 | -83.0587
5.7026 | -83.0592
5.703 | -83.0597
5.7033 | -83.0602
5.7036 | -83.0607
5.704 | -83.0613
5.7043 | -83.0626
5.7046 | -83.0629
5.7049 | -83.0632
5.7053 | -83.0638
5.7056 | -83.0646
5.7059 | -83.065
5.7063 | -83.0655
5.7066 | -83.0665
5.7069 | -83.0668
5.7073 | -83.0675
5.7076 | -83.0686
5.7079 | -83.0702
5.7082 | -83.0707
5.7086 | -83.0712
5.7089 | -83.0716
5.7092 | -83.0723
5.7096 | -83.0729
5.7099 | -83.0733
5.7102 | -83.0736
5.7105 | -83.0736
5.7109 | -83.0737
5.7112 | -83.0744
5.7115 | -83.0748
5.7119 | -83.0756
5.7122 | -83.0764
5.7125 | -83.0767
5.7128 | -83.0772
5.7132 | -83.0782
5.7135 | -83.0789
5.7138 | -83.0803
5.7142 | -83.0829
5.7145 | -83.0843
5.7148 | -83.0847
5.7152 | -83.0854
5.7155 | -83.086
5.7158 | -83.0867
5.7161 | -83.087
5.7165 | -83.0875
5.7168 | -83.0888
5.7171 | -83.0901
5.7175 | -83.0911
5.7178 | -83.0921
5.7181 | -83.0927
5.7184 | -83.0933
5.7188 | -83.0939
5.7191 | -83.0945
5.7194 | -83.0951
5.7198 | -83.0957
5.7201 | -83.0954
5.7204 | -83.0957
5.7207 | -83.0963
5.7211 | -83.0966
5.7214 | -83.097
5.7217 | -83.0969
5.7221 | -83.0969
5.7224 | -83.0973
5.7227 | -83.099
5.723 | -83.0998
5.7234 | -83.1
5.7237 | -83.1004
5.724 | -83.0995
5.7244 | -83.0992
5.7247 | -83.101
5.725 | -83.1044
5.7254 | -83.1071
5.7257 | -83.1079
5.726 | -83.1088
5.7263 | -83.11
5.7267 | -83.111
5.727 | -83.1118
5.7273 | -83.113
5.7277 | -83.1146
5.728 | -83.1155
5.7283 | -83.1162
5.7286 | -83.1164
5.729 | -83.1159
5.7293 | -83.1158
5.7296 | -83.1161
5.73 | -83.1169
5.7303 | -83.1178
5.7306 | -83.1185
5.7309 | -83.1195
5.7313 | -83.1209
5.7316 | -83.1224
5.7319 | -83.1229
5.7323 | -83.1231
5.7326 | -83.1245
5.7329 | -83.1258
5.7333 | -83.1266
5.7336 | -83.1272
5.7339 | -83.1278
5.7342 | -83.1286
5.7346 | -83.1294
5.7349 | -83.1296
5.7352 | -83.1297
5.7356 | -83.1299
5.7359 | -83.1315
5.7362 | -83.1331
5.7365 | -83.1335
5.7369 | -83.1341
5.7372 | -83.1354
5.7375 | -83.1364
5.7379 | -83.1371
5.7382 | -83.138
5.7385 | -83.1388
5.7388 | -83.1392
5.7392 | -83.1392
5.7395 | -83.1397
5.7398 | -83.141
5.7402 | -83.1419
5.7405 | -83.1418
5.7408 | -83.1413
5.7412 | -83.1423
5.7415 | -83.1437
5.7418 | -83.1444
5.7421 | -83.1453
5.7425 | -83.1456
5.7428 | -83.1469
5.7431 | -83.1482
5.7435 | -83.1495
5.7438 | -83.1502
5.7441 | -83.1508
5.7444 | -83.1512
5.7448 | -83.1519
5.7451 | -83.1531
5.7454 | -83.1539
5.7458 | -83.1547
5.7461 | -83.155
5.7464 | -83.156
5.7467 | -83.1578
5.7471 | -83.1585
5.7474 | -83.159
5.7477 | -83.1595
5.7481 | -83.1615
5.7484 | -83.1632
5.7487 | -83.163
5.7491 | -83.1629
5.7494 | -83.1631
5.7497 | -83.1619
5.75 | -83.1616
5.7504 | -83.1624
5.7507 | -83.164
5.751 | -83.1644
5.7514 | -83.1641
5.7517 | -83.1649
5.752 | -83.165
5.7523 | -83.1654
5.7527 | -83.1662
5.753 | -83.1667
5.7533 | -83.1677
5.7537 | -83.1692
5.754 | -83.1707
5.7543 | -83.1715
5.7546 | -83.1723
5.755 | -83.173
5.7553 | -83.1748
5.7556 | -83.1759
5.756 | -83.177
5.7563 | -83.1786
5.7566 | -83.1798
5.757 | -83.1802
5.7573 | -83.1818
5.7576 | -83.183
5.7579 | -83.1838
5.7583 | -83.1843
5.7586 | -83.1849
5.7589 | -83.1857
5.7593 | -83.1865
5.7596 | -83.1864
5.7599 | -83.1864
5.7602 | -83.1869
5.7606 | -83.1876
5.7609 | -83.1885
5.7612 | -83.1894
5.7616 | -83.1905
5.7619 | -83.1911
5.7622 | -83.192
5.7625 | -83.1928
5.7629 | -83.1927
5.7632 | -83.1932
5.7635 | -83.1945
5.7639 | -83.1956
5.7642 | -83.1969
5.7645 | -83.1971
5.7649 | -83.1975
5.7652 | -83.1983
5.7655 | -83.1998
5.7658 | -83.2013
5.7662 | -83.2019
5.7665 | -83.2025
5.7668 | -83.2032
5.7672 | -83.204
5.7675 | -83.205
5.7678 | -83.2059
5.7681 | -83.2061
5.7685 | -83.2049
5.7688 | -83.2054
5.7691 | -83.2065
5.7695 | -83.2071
5.7698 | -83.2078
5.7701 | -83.2085
5.7704 | -83.209
5.7708 | -83.2096
5.7711 | -83.2102
5.7714 | -83.2108
5.7718 | -83.2108
5.7721 | -83.2111
5.7724 | -83.213
5.7728 | -83.2145
5.7731 | -83.2163
5.7734 | -83.2179
5.7737 | -83.2189
5.7741 | -83.2204
5.7744 | -83.2209
5.7747 | -83.2208
5.7751 | -83.2212
5.7754 | -83.2217
5.7757 | -83.2221
5.776 | -83.2235
5.7764 | -83.2261
5.7767 | -83.2264
5.777 | -83.2273
5.7774 | -83.2285
5.7777 | -83.2293
5.778 | -83.2302
5.7783 | -83.2309
5.7787 | -83.2316
5.779 | -83.2327
5.7793 | -83.2342
5.7797 | -83.2352
5.78 | -83.2363
5.7803 | -83.2371
5.7806 | -83.2377
5.781 | -83.2378
5.7813 | -83.2382
5.7816 | -83.2386
5.782 | -83.2391
5.7823 | -83.2398
5.7826 | -83.2408
5.783 | -83.2417
5.7833 | -83.2421
5.7836 | -83.2427
5.7839 | -83.2433
5.7843 | -83.244
5.7846 | -83.2458
5.7849 | -83.2482
5.7853 | -83.2492
5.7856 | -83.2497
5.7859 | -83.251
5.7862 | -83.2531
5.7866 | -83.2544
5.7869 | -83.2551
5.7872 | -83.2561
5.7876 | -83.2568
5.7879 | -83.2571
5.7882 | -83.2576
5.7885 | -83.2583
5.7889 | -83.2599
5.7892 | -83.2621
5.7895 | -83.2633
5.7899 | -83.2638
5.7902 | -83.2643
5.7905 | -83.265
5.7909 | -83.2655
5.7912 | -83.2652
5.7915 | -83.2652
5.7918 | -83.2658
5.7922 | -83.2665
5.7925 | -83.2681
5.7928 | -83.2686
5.7932 | -83.2692
5.7935 | -83.2695
5.7938 | -83.27
5.7941 | -83.2711
5.7945 | -83.2721
5.7948 | -83.2735
5.7951 | -83.2745
5.7955 | -83.2748
5.7958 | -83.2746
5.7961 | -83.2754
5.7964 | -83.276
5.7968 | -83.2774
5.7971 | -83.2788
5.7974 | -83.2797
5.7978 | -83.2807
5.7981 | -83.2815
5.7984 | -83.2823
5.7988 | -83.2832
5.7991 | -83.2842
5.7994 | -83.2849
5.7997 | -83.286
5.8001 | -83.2878
5.8004 | -83.2884
5.8007 | -83.2885
5.8011 | -83.2895
5.8014 | -83.29
5.8017 | -83.2917
5.802 | -83.2931
5.8024 | -83.2951
5.8027 | -83.2957
5.803 | -83.2968
5.8034 | -83.2977
5.8037 | -83.2982
5.804 | -83.2994
5.8043 | -83.3001
5.8047 | -83.3013
5.805 | -83.3024
5.8053 | -83.303
5.8057 | -83.3041
5.806 | -83.305
5.8063 | -83.3056
5.8067 | -83.3064
5.807 | -83.3079
5.8073 | -83.3098
5.8076 | -83.3112
5.808 | -83.3123
5.8083 | -83.3135
5.8086 | -83.3139
5.809 | -83.315
5.8093 | -83.3166
5.8096 | -83.3187
5.8099 | -83.3192
5.8103 | -83.3197
5.8106 | -83.3211
5.8109 | -83.3221
5.8113 | -83.3229
5.8116 | -83.3243
5.8119 | -83.3249
5.8122 | -83.326
5.8126 | -83.3283
5.8129 | -83.3301
5.8132 | -83.331
5.8136 | -83.332
5.8139 | -83.3332
5.8142 | -83.3344
5.8146 | -83.3353
5.8149 | -83.3362
5.8152 | -83.3378
5.8155 | -83.3393
5.8159 | -83.3406
5.8162 | -83.3414
5.8165 | -83.3424
5.8169 | -83.3433
5.8172 | -83.3441
5.8175 | -83.3453
5.8178 | -83.3472
5.8182 | -83.3479
5.8185 | -83.3481
5.8188 | -83.3484
5.8192 | -83.3484
5.8195 | -83.3487
5.8198 | -83.3497
5.8201 | -83.3507
5.8205 | -83.3521
5.8208 | -83.3535
5.8211 | -83.3551
5.8215 | -83.3567
5.8218 | -83.3586
5.8221 | -83.3596
5.8224 | -83.3606
5.8228 | -83.3612
5.8231 | -83.362
5.8234 | -83.363
5.8238 | -83.3639
5.8241 | -83.3648
5.8244 | -83.3658
5.8248 | -83.367
5.8251 | -83.3683
5.8254 | -83.3693
5.8257 | -83.3703
5.8261 | -83.3702
5.8264 | -83.3704
5.8267 | -83.3719
5.8271 | -83.3737
5.8274 | -83.3756
5.8277 | -83.3767
5.828 | -83.3776
5.8284 | -83.3784
5.8287 | -83.3795
5.829 | -83.3801
5.8294 | -83.381
5.8297 | -83.3821
5.83 | -83.3826
5.8303 | -83.3837
5.8307 | -83.385
5.831 | -83.3855
5.8313 | -83.3866
5.8317 | -83.3877
5.832 | -83.3888
5.8323 | -83.3898
5.8327 | -83.3914
5.833 | -83.393
5.8333 | -83.3941
5.8336 | -83.3953
5.834 | -83.3955
5.8343 | -83.3946
5.8346 | -83.3954
5.835 | -83.3963
5.8353 | -83.3978
5.8356 | -83.3996
5.8359 | -83.4006
5.8363 | -83.4018
5.8366 | -83.4024
5.8369 | -83.4022
5.8373 | -83.4028
5.8376 | -83.404
5.8379 | -83.4056
5.8382 | -83.4079
5.8386 | -83.4093
5.8389 | -83.4099
5.8392 | -83.4111
5.8396 | -83.4122
5.8399 | -83.4133
5.8402 | -83.4151
5.8406 | -83.4162
5.8409 | -83.4169
5.8412 | -83.418
5.8415 | -83.4191
5.8419 | -83.4203
5.8422 | -83.4216
5.8425 | -83.4232
5.8429 | -83.4248
5.8432 | -83.4258
5.8435 | -83.4269
5.8438 | -83.4288
5.8442 | -83.4313
5.8445 | -83.4331
5.8448 | -83.4352
5.8452 | -83.4378
5.8455 | -83.44
5.8458 | -83.4407
5.8461 | -83.4414
5.8465 | -83.4426
5.8468 | -83.4441
5.8471 | -83.4454
5.8475 | -83.4464
5.8478 | -83.4473
5.8481 | -83.448
5.8485 | -83.4486
5.8488 | -83.45
5.8491 | -83.4507
5.8494 | -83.4515
5.8498 | -83.4528
5.8501 | -83.4548
5.8504 | -83.4561
5.8508 | -83.4565
5.8511 | -83.4577
5.8514 | -83.4589
5.8517 | -83.4606
5.8521 | -83.462
5.8524 | -83.4633
5.8527 | -83.4649
5.8531 | -83.4659
5.8534 | -83.4675
5.8537 | -83.47
5.854 | -83.4714
5.8544 | -83.4729
5.8547 | -83.4742
5.855 | -83.4745
5.8554 | -83.4773
5.8557 | -83.4803
5.856 | -83.4811
5.8564 | -83.4819
5.8567 | -83.4831
5.857 | -83.4845
5.8573 | -83.4859
5.8577 | -83.4876
5.858 | -83.4888
5.8583 | -83.4904
5.8587 | -83.4917
5.859 | -83.493
5.8593 | -83.494
5.8596 | -83.4953
5.86 | -83.4967
5.8603 | -83.4977
5.8606 | -83.4987
5.861 | -83.5001
5.8613 | -83.5015
5.8616 | -83.5031
5.8619 | -83.5044
5.8623 | -83.5051
5.8626 | -83.5061
5.8629 | -83.5073
5.8633 | -83.509
5.8636 | -83.5097
5.8639 | -83.5107
5.8643 | -83.5112
5.8646 | -83.5127
5.8649 | -83.5148
5.8652 | -83.5166
5.8656 | -83.5174
5.8659 | -83.519
5.8662 | -83.5208
5.8666 | -83.522
5.8669 | -83.5228
5.8672 | -83.5222
5.8675 | -83.5227
5.8679 | -83.5238
5.8682 | -83.525
5.8685 | -83.5266
5.8689 | -83.5275
5.8692 | -83.5283
5.8695 | -83.5284
5.8698 | -83.5293
5.8702 | -83.5308
5.8705 | -83.5321
5.8708 | -83.5332
5.8712 | -83.5344
5.8715 | -83.5373
5.8718 | -83.5405
5.8722 | -83.5421
5.8725 | -83.5432
5.8728 | -83.5447
5.8731 | -83.5451
5.8735 | -83.5461
5.8738 | -83.5478
5.8741 | -83.5493
5.8745 | -83.55
5.8748 | -83.5504
5.8751 | -83.5507
5.8754 | -83.5519
5.8758 | -83.5531
5.8761 | -83.554
5.8764 | -83.5559
5.8768 | -83.5582
5.8771 | -83.5596
5.8774 | -83.5606
5.8777 | -83.5616
5.8781 | -83.5632
5.8784 | -83.5647
5.8787 | -83.566
5.8791 | -83.5669
5.8794 | -83.5682
5.8797 | -83.5696
5.88 | -83.5706
5.8804 | -83.572
5.8807 | -83.5733
5.881 | -83.5749
5.8814 | -83.5757
5.8817 | -83.5764
5.882 | -83.5775
5.8824 | -83.5789
5.8827 | -83.5811
5.883 | -83.5823
5.8833 | -83.5836
5.8837 | -83.5847
5.884 | -83.5848
5.8843 | -83.5854
5.8847 | -83.5866
5.885 | -83.5878
5.8853 | -83.5891
5.8856 | -83.5902
5.886 | -83.5897
5.8863 | -83.5893
5.8866 | -83.5904
5.887 | -83.5911
5.8873 | -83.5919
5.8876 | -83.5937
5.8879 | -83.5939
5.8883 | -83.5938
5.8886 | -83.5945
5.8889 | -83.5965
5.8893 | -83.598
5.8896 | -83.5992
5.8899 | -83.6013
5.8903 | -83.602
5.8906 | -83.6029
5.8909 | -83.6037
5.8912 | -83.6045
5.8916 | -83.6052
5.8919 | -83.6053
5.8922 | -83.6071
5.8926 | -83.6088
5.8929 | -83.6097
5.8932 | -83.6108
5.8935 | -83.6121
5.8939 | -83.6127
5.8942 | -83.6125
5.8945 | -83.6139
5.8949 | -83.6153
5.8952 | -83.6167
5.8955 | -83.6177
5.8958 | -83.6188
5.8962 | -83.6198
5.8965 | -83.6198
5.8968 | -83.6201
5.8972 | -83.6198
5.8975 | -83.6211
5.8978 | -83.6224
5.8982 | -83.6231
5.8985 | -83.6244
5.8988 | -83.6257
5.8991 | -83.6267
5.8995 | -83.6276
5.8998 | -83.6275
5.9001 | -83.6274
5.9005 | -83.6289
5.9008 | -83.6305
5.9011 | -83.6308
5.9014 | -83.6315
5.9018 | -83.6335
5.9021 | -83.6361
5.9024 | -83.6373
5.9028 | -83.6386
5.9031 | -83.639
5.9034 | -83.6399
5.9037 | -83.6416
5.9041 | -83.6426
5.9044 | -83.6437
5.9047 | -83.6449
5.9051 | -83.6462
5.9054 | -83.6473
5.9057 | -83.6474
5.9061 | -83.648
5.9064 | -83.6491
5.9067 | -83.6503
5.907 | -83.6509
5.9074 | -83.6507
5.9077 | -83.6509
5.908 | -83.6526
5.9084 | -83.6541
5.9087 | -83.6552
5.909 | -83.6561
5.9093 | -83.6567
5.9097 | -83.6569
5.91 | -83.6584
5.9103 | -83.6596
5.9107 | -83.6602
5.911 | -83.6615
5.9113 | -83.6632
5.9116 | -83.6643
5.912 | -83.6647
5.9123 | -83.6651
5.9126 | -83.666
5.913 | -83.6671
5.9133 | -83.6687
5.9136 | -83.6699
5.9139 | -83.6703
5.9143 | -83.6704
5.9146 | -83.671
5.9149 | -83.6716
5.9153 | -83.6724
5.9156 | -83.6733
5.9159 | -83.6742
5.9163 | -83.6745
5.9166 | -83.6745
5.9169 | -83.6752
5.9172 | -83.6756
5.9176 | -83.6765
5.9179 | -83.678
5.9182 | -83.6804
5.9186 | -83.6816
5.9189 | -83.683
5.9192 | -83.6837
5.9195 | -83.6839
5.9199 | -83.6847
5.9202 | -83.685
5.9205 | -83.685
5.9209 | -83.6851
5.9212 | -83.6854
5.9215 | -83.686
5.9218 | -83.6866
5.9222 | -83.687
5.9225 | -83.6876
5.9228 | -83.6881
5.9232 | -83.6888
5.9235 | -83.6896
5.9238 | -83.6902
5.9242 | -83.6909
5.9245 | -83.6917
5.9248 | -83.6927
5.9251 | -83.6936
5.9255 | -83.6946
5.9258 | -83.6958
5.9261 | -83.6951
5.9265 | -83.695
5.9268 | -83.6958
5.9271 | -83.6961
5.9274 | -83.6956
5.9278 | -83.6958
5.9281 | -83.698
5.9284 | -83.6994
5.9288 | -83.7002
5.9291 | -83.7009
5.9294 | -83.7015
5.9297 | -83.702
5.9301 | -83.702
5.9304 | -83.7023
5.9307 | -83.7032
5.9311 | -83.7046
5.9314 | -83.7052
5.9317 | -83.7064
5.9321 | -83.7078
5.9324 | -83.7093
5.9327 | -83.7105
5.933 | -83.7111
5.9334 | -83.7114
5.9337 | -83.7117
5.934 | -83.7119
5.9344 | -83.7112
5.9347 | -83.7119
5.935 | -83.7123
5.9353 | -83.7138
5.9357 | -83.7158
5.936 | -83.7194
5.9363 | -83.7212
5.9367 | -83.7219
5.937 | -83.7217
5.9373 | -83.7209
5.9376 | -83.7217
5.938 | -83.7221
5.9383 | -83.7222
5.9386 | -83.7233
5.939 | -83.7247
5.9393 | -83.7244
5.9396 | -83.7239
5.94 | -83.7247
5.9403 | -83.7258
5.9406 | -83.7262
5.9409 | -83.7267
5.9413 | -83.7276
5.9416 | -83.7272
5.9419 | -83.7268
5.9423 | -83.7273
5.9426 | -83.729
5.9429 | -83.7296
5.9432 | -83.7295
5.9436 | -83.7301
5.9439 | -83.7305
5.9442 | -83.7326
5.9446 | -83.733
5.9449 | -83.734
5.9452 | -83.7347
5.9455 | -83.7355
5.9459 | -83.7359
5.9462 | -83.7369
5.9465 | -83.7374
5.9469 | -83.7381
5.9472 | -83.7395
5.9475 | -83.7407
5.9479 | -83.7414
5.9482 | -83.7422
5.9485 | -83.7432
5.9488 | -83.744
5.9492 | -83.7441
5.9495 | -83.7447
5.9498 | -83.7451
5.9502 | -83.7458
5.9505 | -83.7473
5.9508 | -83.7494
5.9511 | -83.7512
5.9515 | -83.752
5.9518 | -83.7524
5.9521 | -83.7529
5.9525 | -83.7532
5.9528 | -83.7537
5.9531 | -83.7539
5.9534 | -83.7538
5.9538 | -83.7536
5.9541 | -83.7536
5.9544 | -83.7538
5.9548 | -83.754
5.9551 | -83.7543
5.9554 | -83.7544
5.9558 | -83.7545
5.9561 | -83.7549
5.9564 | -83.7547
5.9567 | -83.7537
5.9571 | -83.7535
5.9574 | -83.7534
5.9577 | -83.7534
5.9581 | -83.754
5.9584 | -83.7541
5.9587 | -83.7542
5.959 | -83.7546
5.9594 | -83.7551
5.9597 | -83.7553
5.96 | -83.755
5.9604 | -83.7549
5.9607 | -83.754
5.961 | -83.7539
5.9613 | -83.7528
5.9617 | -83.7514
5.962 | -83.7508
5.9623 | -83.7511
5.9627 | -83.7509
5.963 | -83.7506
5.9633 | -83.7501
5.9637 | -83.7501
5.964 | -83.7508
5.9643 | -83.7512
5.9646 | -83.751
5.965 | -83.7511
5.9653 | -83.7509
5.9656 | -83.7511
5.966 | -83.7513
5.9663 | -83.7518
5.9666 | -83.7522
5.9669 | -83.7521
5.9673 | -83.752
5.9676 | -83.752
5.9679 | -83.7521
5.9683 | -83.7523
5.9686 | -83.7525
5.9689 | -83.7537
5.9692 | -83.7544
5.9696 | -83.7542
5.9699 | -83.7552
5.9702 | -83.7571
5.9706 | -83.7591
5.9709 | -83.7596
5.9712 | -83.7596
5.9715 | -83.7596
5.9719 | -83.7604
5.9722 | -83.7608
5.9725 | -83.7601
5.9729 | -83.7603
5.9732 | -83.7606
5.9735 | -83.7615
5.9739 | -83.7619
5.9742 | -83.7619
5.9745 | -83.7621
5.9748 | -83.7622
5.9752 | -83.763
5.9755 | -83.7635
5.9758 | -83.7631
5.9762 | -83.7629
5.9765 | -83.7631
5.9768 | -83.7633
5.9771 | -83.7635
5.9775 | -83.7638
5.9778 | -83.7644
5.9781 | -83.7641
5.9785 | -83.7643
5.9788 | -83.7647
5.9791 | -83.7648
5.9794 | -83.7644
5.9798 | -83.7651
5.9801 | -83.7658
5.9804 | -83.7665
5.9808 | -83.7669
5.9811 | -83.7667
5.9814 | -83.7667
5.9818 | -83.7666
5.9821 | -83.7667
5.9824 | -83.7678
5.9827 | -83.7681
5.9831 | -83.7681
5.9834 | -83.7691
5.9837 | -83.771
5.9841 | -83.7723
5.9844 | -83.7728
5.9847 | -83.7727
5.985 | -83.7728
5.9854 | -83.7729
5.9857 | -83.7732
5.986 | -83.7733
5.9864 | -83.7729
5.9867 | -83.7727
5.987 | -83.7724
5.9873 | -83.7722
5.9877 | -83.7724
5.988 | -83.7726
5.9883 | -83.7731
5.9887 | -83.7741
5.989 | -83.7738
5.9893 | -83.7731
5.9897 | -83.7735
5.99 | -83.7736
5.9903 | -83.7737
5.9906 | -83.7735
5.991 | -83.7737
5.9913 | -83.7737
5.9916 | -83.7734
5.992 | -83.7698
5.9923 | -83.7697
5.9926 | -83.7689
5.9929 | -83.7688
5.9933 | -83.7695
5.9936 | -83.77
5.9939 | -83.7702
5.9943 | -83.7702
5.9946 | -83.7701
5.9949 | -83.7699
5.9952 | -83.7695
5.9956 | -83.7692
5.9959 | -83.7693
5.9962 | -83.7692
5.9966 | -83.7705
5.9969 | -83.7719
5.9972 | -83.772
5.9976 | -83.7719
5.9979 | -83.7719
5.9982 | -83.7718
5.9985 | -83.7721
5.9989 | -83.7719
5.9992 | -83.7715
5.9995 | -83.7715
5.9999 | -83.7713
6.0002 | -83.7706
6.0005 | -83.7699
6.0008 | -83.7692
6.0012 | -83.7688
6.0015 | -83.7686
6.0018 | -83.7686
6.0022 | -83.7698
6.0025 | -83.7706
6.0028 | -83.7705
6.0031 | -83.7707
6.0035 | -83.7713
6.0038 | -83.7711
6.0041 | -83.7714
6.0045 | -83.7721
6.0048 | -83.7723
6.0051 | -83.7725
6.0055 | -83.7728
6.0058 | -83.7739
6.0061 | -83.7743
6.0064 | -83.7739
6.0068 | -83.7737
6.0071 | -83.7736
6.0074 | -83.7734
6.0078 | -83.7731
6.0081 | -83.7728
6.0084 | -83.7732
6.0087 | -83.7741
6.0091 | -83.7743
6.0094 | -83.7746
6.0097 | -83.7755
6.0101 | -83.7758
6.0104 | -83.7762
6.0107 | -83.7766
6.011 | -83.7751
6.0114 | -83.7746
6.0117 | -83.7744
6.012 | -83.7741
6.0124 | -83.7738
6.0127 | -83.7732
6.013 | -83.7724
6.0133 | -83.7728
6.0137 | -83.7724
6.014 | -83.7719
6.0143 | -83.7717
6.0147 | -83.7725
6.015 | -83.7732
6.0153 | -83.773
6.0157 | -83.7728
6.016 | -83.7725
6.0163 | -83.7729
6.0166 | -83.7726
6.017 | -83.7721
6.0173 | -83.772
6.0176 | -83.7718
6.018 | -83.772
6.0183 | -83.7715
6.0186 | -83.7711
6.0189 | -83.7708
6.0193 | -83.7705
6.0196 | -83.7704
6.0199 | -83.7711
6.0203 | -83.772
6.0206 | -83.7723
6.0209 | -83.7724
6.0212 | -83.7723
6.0216 | -83.7724
6.0219 | -83.772
6.0222 | -83.7715
6.0226 | -83.7709
6.0229 | -83.7707
6.0232 | -83.7699
6.0236 | -83.7696
6.0239 | -83.7708
6.0242 | -83.7718
6.0245 | -83.7706
6.0249 | -83.7699
6.0252 | -83.77
6.0255 | -83.7705
6.0259 | -83.7706
6.0262 | -83.7703
6.0265 | -83.7702
6.0268 | -83.7704
6.0272 | -83.7703
6.0275 | -83.7697
6.0278 | -83.7696
6.0282 | -83.7692
6.0285 | -83.7694
6.0288 | -83.7692
6.0291 | -83.769
6.0295 | -83.7688
6.0298 | -83.7678
6.0301 | -83.7678
6.0305 | -83.7676
6.0308 | -83.7677
6.0311 | -83.767
6.0315 | -83.7658
6.0318 | -83.7655
6.0321 | -83.7661
6.0324 | -83.7662
6.0328 | -83.7659
6.0331 | -83.765
6.0334 | -83.7645
6.0338 | -83.7655
6.0341 | -83.7658
6.0344 | -83.7652
6.0347 | -83.765
6.0351 | -83.7648
6.0354 | -83.7646
6.0357 | -83.765
6.0361 | -83.766
6.0364 | -83.7658
6.0367 | -83.7634
6.037 | -83.7612
6.0374 | -83.7603
6.0377 | -83.7602
6.038 | -83.7587
6.0384 | -83.7578
6.0387 | -83.7576
6.039 | -83.7576
6.0394 | -83.7575
6.0397 | -83.7574
6.04 | -83.7577
6.0403 | -83.7577
6.0407 | -83.7581
6.041 | -83.758
6.0413 | -83.7577
6.0417 | -83.7572
6.042 | -83.7568
6.0423 | -83.7569
6.0426 | -83.7569
6.043 | -83.7558
6.0433 | -83.7549
6.0436 | -83.7552
6.044 | -83.7555
6.0443 | -83.7552
6.0446 | -83.7552
6.0449 | -83.7557
6.0453 | -83.7552
6.0456 | -83.7549
6.0459 | -83.7549
6.0463 | -83.7548
6.0466 | -83.7551
6.0469 | -83.756
6.0473 | -83.7547
6.0476 | -83.7525
6.0479 | -83.7506
6.0482 | -83.7495
6.0486 | -83.749
6.0489 | -83.7486
6.0492 | -83.7485
6.0496 | -83.7481
6.0499 | -83.7483
6.0502 | -83.7479
6.0505 | -83.7476
6.0509 | -83.7472
6.0512 | -83.7472
6.0515 | -83.7474
6.0519 | -83.7476
6.0522 | -83.7484
6.0525 | -83.7493
6.0528 | -83.7492
6.0532 | -83.7488
6.0535 | -83.7487
6.0538 | -83.748
6.0542 | -83.748
6.0545 | -83.7477
6.0548 | -83.7468
6.0552 | -83.746
6.0555 | -83.7458
6.0558 | -83.7457
6.0561 | -83.7461
6.0565 | -83.746
6.0568 | -83.7457
6.0571 | -83.7455
6.0575 | -83.7453
6.0578 | -83.7455
6.0581 | -83.7457
6.0584 | -83.7454
6.0588 | -83.7455
6.0591 | -83.7455
6.0594 | -83.746
6.0598 | -83.7464
6.0601 | -83.7461
6.0604 | -83.7457
6.0607 | -83.7452
6.0611 | -83.7448
6.0614 | -83.7441
6.0617 | -83.7437
6.0621 | -83.7434
6.0624 | -83.7435
6.0627 | -83.7439
6.063 | -83.7439
6.0634 | -83.7432
6.0637 | -83.7431
6.064 | -83.7441
6.0644 | -83.7445
6.0647 | -83.7441
6.065 | -83.7438
6.0654 | -83.7436
6.0657 | -83.7437
6.066 | -83.7446
6.0663 | -83.7453
6.0667 | -83.7461
6.067 | -83.7464
6.0673 | -83.7459
6.0677 | -83.7455
6.068 | -83.7463
6.0683 | -83.7462
6.0686 | -83.7455
6.069 | -83.7435
6.0693 | -83.7419
6.0696 | -83.7422
6.07 | -83.7435
6.0703 | -83.7435
6.0706 | -83.7435
6.0709 | -83.743
6.0713 | -83.7431
6.0716 | -83.7436
6.0719 | -83.7438
6.0723 | -83.7439
6.0726 | -83.7438
6.0729 | -83.7438
6.0733 | -83.7435
6.0736 | -83.7435
6.0739 | -83.7435
6.0742 | -83.7433
6.0746 | -83.743
6.0749 | -83.7422
6.0752 | -83.7409
6.0756 | -83.7395
6.0759 | -83.7385
6.0762 | -83.7383
6.0765 | -83.7384
6.0769 | -83.7389
6.0772 | -83.7389
6.0775 | -83.7391
6.0779 | -83.738
6.0782 | -83.7374
6.0785 | -83.7368
6.0788 | -83.7364
6.0792 | -83.7361
6.0795 | -83.736
6.0798 | -83.7364
6.0802 | -83.7371
6.0805 | -83.736
6.0808 | -83.735
6.0812 | -83.7348
6.0815 | -83.7348
6.0818 | -83.7347
6.0821 | -83.7352
6.0825 | -83.7354
6.0828 | -83.7355
6.0831 | -83.7352
6.0835 | -83.7337
6.0838 | -83.7328
6.0841 | -83.7334
6.0844 | -83.7352
6.0848 | -83.7361
6.0851 | -83.7358
6.0854 | -83.7354
6.0858 | -83.7352
6.0861 | -83.735
6.0864 | -83.736
6.0867 | -83.736
6.0871 | -83.7359
6.0874 | -83.7358
6.0877 | -83.7367
6.0881 | -83.7371
6.0884 | -83.739
6.0887 | -83.74
6.0891 | -83.7397
6.0894 | -83.7395
6.0897 | -83.7396
6.09 | -83.7402
6.0904 | -83.7402
6.0907 | -83.7401
6.091 | -83.7404
6.0914 | -83.7426
6.0917 | -83.7444
6.092 | -83.7448
6.0923 | -83.7445
6.0927 | -83.7452
6.093 | -83.7458
6.0933 | -83.7459
6.0937 | -83.746
6.094 | -83.7462
6.0943 | -83.7462
6.0946 | -83.7461
6.095 | -83.7462
6.0953 | -83.7458
6.0956 | -83.7452
6.096 | -83.7454
6.0963 | -83.7462
6.0966 | -83.7471
6.097 | -83.7481
6.0973 | -83.7492
6.0976 | -83.75
6.0979 | -83.7501
6.0983 | -83.7503
6.0986 | -83.7504
6.0989 | -83.7506
6.0993 | -83.7509
6.0996 | -83.7507
6.0999 | -83.751
6.1002 | -83.7514
6.1006 | -83.752
6.1009 | -83.7523
6.1012 | -83.7516
6.1016 | -83.7504
6.1019 | -83.7501
6.1022 | -83.7494
6.1025 | -83.7497
6.1029 | -83.7494
6.1032 | -83.7493
6.1035 | -83.7496
6.1039 | -83.75
6.1042 | -83.7497
6.1045 | -83.7484
6.1048 | -83.7487
6.1052 | -83.7494
6.1055 | -83.7497
6.1058 | -83.7501
6.1062 | -83.7513
6.1065 | -83.7518
6.1068 | -83.7517
6.1072 | -83.752
6.1075 | -83.7525
6.1078 | -83.7527
6.1081 | -83.7529
6.1085 | -83.7531
6.1088 | -83.7533
6.1091 | -83.7534
6.1095 | -83.7535
6.1098 | -83.7536
6.1101 | -83.7535
6.1104 | -83.7538
6.1108 | -83.7541
6.1111 | -83.7532
6.1114 | -83.7532
6.1118 | -83.7533
6.1121 | -83.7534
6.1124 | -83.7531
6.1127 | -83.7526
6.1131 | -83.7531
6.1134 | -83.7536
6.1137 | -83.7537
6.1141 | -83.7542
6.1144 | -83.754
6.1147 | -83.7541
6.1151 | -83.7539
6.1154 | -83.7547
6.1157 | -83.7547
6.116 | -83.7545
6.1164 | -83.7548
6.1167 | -83.7549
6.117 | -83.7553
6.1174 | -83.7558
6.1177 | -83.7563
6.118 | -83.7564
6.1183 | -83.7569
6.1187 | -83.757
6.119 | -83.7574
6.1193 | -83.7576
6.1197 | -83.7569
6.12 | -83.7549
6.1203 | -83.7546
6.1206 | -83.7551
6.121 | -83.756
6.1213 | -83.7571
6.1216 | -83.7581
6.122 | -83.7584
6.1223 | -83.7589
6.1226 | -83.7593
6.123 | -83.7592
6.1233 | -83.7593
6.1236 | -83.7598
6.1239 | -83.7602
6.1243 | -83.7608
6.1246 | -83.7614
6.1249 | -83.7627
6.1253 | -83.7631
6.1256 | -83.7634
6.1259 | -83.7638
6.1262 | -83.7642
6.1266 | -83.7646
6.1269 | -83.765
6.1272 | -83.7659
6.1276 | -83.7666
6.1279 | -83.767
6.1282 | -83.7674
6.1285 | -83.7678
6.1289 | -83.7683
6.1292 | -83.7686
6.1295 | -83.769
6.1299 | -83.7692
6.1302 | -83.7696
6.1305 | -83.7701
6.1309 | -83.7702
6.1312 | -83.7701
6.1315 | -83.7718
6.1318 | -83.7736
6.1322 | -83.7744
6.1325 | -83.7754
6.1328 | -83.7762
6.1332 | -83.7772
6.1335 | -83.7788
6.1338 | -83.7798
6.1341 | -83.7802
6.1345 | -83.7809
6.1348 | -83.7812
6.1351 | -83.7818
6.1355 | -83.7822
6.1358 | -83.7842
6.1361 | -83.7858
6.1364 | -83.7861
6.1368 | -83.7873
6.1371 | -83.7884
6.1374 | -83.7889
6.1378 | -83.7892
6.1381 | -83.7898
6.1384 | -83.7905
6.1388 | -83.7909
6.1391 | -83.7914
6.1394 | -83.7919
6.1397 | -83.7918
6.1401 | -83.7921
6.1404 | -83.7928
6.1407 | -83.7933
6.1411 | -83.7951
6.1414 | -83.7965
6.1417 | -83.7969
6.142 | -83.7967
6.1424 | -83.7969
6.1427 | -83.7976
6.143 | -83.7983
6.1434 | -83.799
6.1437 | -83.7989
6.144 | -83.7998
6.1443 | -83.8004
6.1447 | -83.8009
6.145 | -83.801
6.1453 | -83.8018
6.1457 | -83.8026
6.146 | -83.8029
6.1463 | -83.8034
6.1467 | -83.8044
6.147 | -83.8052
6.1473 | -83.8061
6.1476 | -83.8071
6.148 | -83.8079
6.1483 | -83.8093
6.1486 | -83.8099
6.149 | -83.8108
6.1493 | -83.8116
6.1496 | -83.8122
6.1499 | -83.813
6.1503 | -83.8139
6.1506 | -83.8148
6.1509 | -83.8158
6.1513 | -83.8155
6.1516 | -83.8154
6.1519 | -83.8169
6.1522 | -83.8181
6.1526 | -83.8188
6.1529 | -83.8195
6.1532 | -83.8206
6.1536 | -83.8221
6.1539 | -83.8235
6.1542 | -83.8241
6.1546 | -83.8255
6.1549 | -83.8277
6.1552 | -83.8286
6.1555 | -83.8292
6.1559 | -83.8293
6.1562 | -83.8299
6.1565 | -83.8306
6.1569 | -83.8312
6.1572 | -83.831
6.1575 | -83.8316
6.1578 | -83.8319
6.1582 | -83.8315
6.1585 | -83.8319
6.1588 | -83.8327
6.1592 | -83.8327
6.1595 | -83.833
6.1598 | -83.8332
6.1601 | -83.8336
6.1605 | -83.8342
6.1608 | -83.8348
6.1611 | -83.8358
6.1615 | -83.8363
6.1618 | -83.8376
6.1621 | -83.8391
6.1624 | -83.8409
6.1628 | -83.8425
6.1631 | -83.8433
6.1634 | -83.8452
6.1638 | -83.8465
6.1641 | -83.8468
6.1644 | -83.8464
6.1648 | -83.8468
6.1651 | -83.8483
6.1654 | -83.8491
6.1657 | -83.8504
6.1661 | -83.8509
6.1664 | -83.851
6.1667 | -83.8516
6.1671 | -83.8525
6.1674 | -83.8531
6.1677 | -83.8538
6.168 | -83.8545
6.1684 | -83.8551
6.1687 | -83.8559
6.169 | -83.8561
6.1694 | -83.8568
6.1697 | -83.8576
6.17 | -83.8582
6.1703 | -83.8585
6.1707 | -83.8592
6.171 | -83.8602
6.1713 | -83.8616
6.1717 | -83.8626
6.172 | -83.8634
6.1723 | -83.8644
6.1727 | -83.8654
6.173 | -83.8664
6.1733 | -83.8668
6.1736 | -83.8664
6.174 | -83.8669
6.1743 | -83.8679
6.1746 | -83.8689
6.175 | -83.8691
6.1753 | -83.8685
6.1756 | -83.8694
6.1759 | -83.8703
6.1763 | -83.871
6.1766 | -83.8717
6.1769 | -83.872
6.1773 | -83.8731
6.1776 | -83.874
6.1779 | -83.8749
6.1782 | -83.8756
6.1786 | -83.8765
6.1789 | -83.8775
6.1792 | -83.8787
6.1796 | -83.88
6.1799 | -83.881
6.1802 | -83.8822
6.1806 | -83.883
6.1809 | -83.8834
6.1812 | -83.8843
6.1815 | -83.8853
6.1819 | -83.8855
6.1822 | -83.8851
6.1825 | -83.8857
6.1829 | -83.886
6.1832 | -83.886
6.1835 | -83.8888
6.1838 | -83.8905
6.1842 | -83.8906
6.1845 | -83.8912
6.1848 | -83.8921
6.1852 | -83.8929
6.1855 | -83.8938
6.1858 | -83.8942
6.1861 | -83.8946
6.1865 | -83.8961
6.1868 | -83.8972
6.1871 | -83.8988
6.1875 | -83.8996
6.1878 | -83.9005
6.1881 | -83.9013
6.1885 | -83.9022
6.1888 | -83.9019
6.1891 | -83.9006
6.1894 | -83.9006
6.1898 | -83.9017
6.1901 | -83.9035
6.1904 | -83.9045
6.1908 | -83.9054
6.1911 | -83.9063
6.1914 | -83.9071
6.1917 | -83.9064
6.1921 | -83.9068
6.1924 | -83.9084
6.1927 | -83.909
6.1931 | -83.9092
6.1934 | -83.91
6.1937 | -83.9101
6.194 | -83.9106
6.1944 | -83.9111
6.1947 | -83.9123
6.195 | -83.9131
6.1954 | -83.9137
6.1957 | -83.9142
6.196 | -83.9147
6.1964 | -83.916
6.1967 | -83.9169
6.197 | -83.9168
6.1973 | -83.9178
6.1977 | -83.9184
6.198 | -83.9182
6.1983 | -83.9186
6.1987 | -83.9195
6.199 | -83.9215
6.1993 | -83.9223
6.1996 | -83.9226
6.2 | -83.9233
6.2003 | -83.9236
6.2006 | -83.9245
6.201 | -83.9259
6.2013 | -83.927
6.2016 | -83.9277
6.2019 | -83.929
6.2023 | -83.9301
6.2026 | -83.9307
6.2029 | -83.9314
6.2033 | -83.9311
6.2036 | -83.9311
6.2039 | -83.9319
6.2043 | -83.9326
6.2046 | -83.9326
6.2049 | -83.9335
6.2052 | -83.9341
6.2056 | -83.9346
6.2059 | -83.9353
6.2062 | -83.935
6.2066 | -83.935
6.2069 | -83.9354
6.2072 | -83.9358
6.2075 | -83.9373
6.2079 | -83.9384
6.2082 | -83.9391
6.2085 | -83.9392
6.2089 | -83.9399
6.2092 | -83.9407
6.2095 | -83.9413
6.2098 | -83.942
6.2102 | -83.9435
6.2105 | -83.9438
6.2108 | -83.9457
6.2112 | -83.9466
6.2115 | -83.9477
6.2118 | -83.9486
6.2121 | -83.9489
6.2125 | -83.9497
6.2128 | -83.9516
6.2131 | -83.9523
6.2135 | -83.9532
6.2138 | -83.9536
6.2141 | -83.954
6.2145 | -83.9547
6.2148 | -83.9556
6.2151 | -83.9565
6.2154 | -83.9569
6.2158 | -83.9572
6.2161 | -83.957
6.2164 | -83.9574
6.2168 | -83.9572
6.2171 | -83.9556
6.2174 | -83.9557
6.2177 | -83.9574
6.2181 | -83.9583
6.2184 | -83.9599
6.2187 | -83.9615
6.2191 | -83.9624
6.2194 | -83.9632
6.2197 | -83.9639
6.22 | -83.9646
6.2204 | -83.9654
6.2207 | -83.9664
6.221 | -83.9671
6.2214 | -83.9679
6.2217 | -83.9683
6.222 | -83.9691
6.2224 | -83.9692
6.2227 | -83.9698
6.223 | -83.9709
6.2233 | -83.9711
6.2237 | -83.9712
6.224 | -83.9722
6.2243 | -83.9735
6.2247 | -83.9744
6.225 | -83.9747
6.2253 | -83.9743
6.2256 | -83.9743
6.226 | -83.9751
6.2263 | -83.9762
6.2266 | -83.9771
6.227 | -83.9775
6.2273 | -83.9782
6.2276 | -83.9792
6.2279 | -83.9803
6.2283 | -83.9807
6.2286 | -83.981
6.2289 | -83.9821
6.2293 | -83.9823
6.2296 | -83.9828
6.2299 | -83.9833
6.2303 | -83.9845
6.2306 | -83.9858
6.2309 | -83.9858
6.2312 | -83.984
6.2316 | -83.9835
6.2319 | -83.9837
6.2322 | -83.9843
6.2326 | -83.985
6.2329 | -83.9852
6.2332 | -83.9851
6.2335 | -83.9854
6.2339 | -83.9859
6.2342 | -83.9861
6.2345 | -83.9861
6.2349 | -83.9864
6.2352 | -83.9869
6.2355 | -83.9875
6.2358 | -83.9878
6.2362 | -83.9876
6.2365 | -83.9877
6.2368 | -83.9877
6.2372 | -83.9882
6.2375 | -83.9899
6.2378 | -83.9907
6.2382 | -83.9907
6.2385 | -83.9905
6.2388 | -83.9908
6.2391 | -83.9909
6.2395 | -83.9911
6.2398 | -83.9914
6.2401 | -83.9919
6.2405 | -83.9922
6.2408 | -83.9927
6.2411 | -83.9932
6.2414 | -83.9933
6.2418 | -83.9942
6.2421 | -83.9948
6.2424 | -83.9944
6.2428 | -83.9944
6.2431 | -83.9945
6.2434 | -83.9944
6.2437 | -83.994
6.2441 | -83.9941
6.2444 | -83.9946
6.2447 | -83.995
6.2451 | -83.9969
6.2454 | -83.9983
6.2457 | -83.9986
6.2461 | -83.9988
6.2464 | -83.9987
6.2467 | -83.9986
6.247 | -83.999
6.2474 | -83.9996
6.2477 | -84
6.248 | -84.0006
6.2484 | -84.0013
6.2487 | -84.0019
6.249 | -84.0019
6.2493 | -84.0026
6.2497 | -84.0031
6.25 | -84.0034
6.2503 | -84.0034
6.2507 | -84.0041
6.251 | -84.0043
6.2513 | -84.0038
6.2516 | -84.0032
6.252 | -84.0036
6.2523 | -84.004
6.2526 | -84.0042
6.253 | -84.0042
6.2533 | -84.0042
6.2536 | -84.0056
6.2539 | -84.0078
6.2543 | -84.008
6.2546 | -84.008
6.2549 | -84.0085
6.2553 | -84.0093
6.2556 | -84.0096
6.2559 | -84.0094
6.2563 | -84.0098
6.2566 | -84.0114
6.2569 | -84.0126
6.2572 | -84.0126
6.2576 | -84.0125
6.2579 | -84.0125
6.2582 | -84.0115
6.2586 | -84.0117
6.2589 | -84.0121
6.2592 | -84.0122
6.2595 | -84.0122
6.2599 | -84.0122
6.2602 | -84.0119
6.2605 | -84.0118
6.2609 | -84.012
6.2612 | -84.0126
6.2615 | -84.0132
6.2618 | -84.0135
6.2622 | -84.0132
6.2625 | -84.0129
6.2628 | -84.013
6.2632 | -84.0128
6.2635 | -84.0127
6.2638 | -84.013
6.2642 | -84.0128
6.2645 | -84.0114
6.2648 | -84.0114
6.2651 | -84.0112
6.2655 | -84.0102
6.2658 | -84.009
6.2661 | -84.0087
6.2665 | -84.0084
6.2668 | -84.0085
6.2671 | -84.0083
6.2674 | -84.0064
6.2678 | -84.0056
6.2681 | -84.0052
6.2684 | -84.0037
6.2688 | -84.0033
6.2691 | -84.004
6.2694 | -84.004
6.2697 | -84.0045
6.2701 | -84.005
6.2704 | -84.0047
6.2707 | -84.0043
6.2711 | -84.0041
6.2714 | -84.0041
6.2717 | -84.0047
6.2721 | -84.0046
6.2724 | -84.0047
6.2727 | -84.0048
6.273 | -84.0043
6.2734 | -84.0036
6.2737 | -84.0033
6.274 | -84.0035
6.2744 | -84.0037
6.2747 | -84.004
6.275 | -84.004
6.2753 | -84.004
6.2757 | -84.0045
6.276 | -84.0057
6.2763 | -84.0058
6.2767 | -84.0054
6.277 | -84.0038
6.2773 | -84.0038
6.2776 | -84.0038
6.278 | -84.0036
6.2783 | -84.0036
6.2786 | -84.0035
6.279 | -84.0039
6.2793 | -84.0043
6.2796 | -84.0046
6.28 | -84.0046
6.2803 | -84.0042
6.2806 | -84.0044
6.2809 | -84.0047
6.2813 | -84.0046
6.2816 | -84.0038
6.2819 | -84.0031
6.2823 | -84.0029
6.2826 | -84.0027
6.2829 | -84.0026
6.2832 | -84.0017
6.2836 | -84.0005
6.2839 | -84.0008
6.2842 | -84.0009
6.2846 | -84.001
6.2849 | -84.001
6.2852 | -84.0014
6.2855 | -84.0015
6.2859 | -84.0017
6.2862 | -84.0014
6.2865 | -84.0014
6.2869 | -84.0007
6.2872 | -84.0004
6.2875 | -84.0008
6.2879 | -84.0019
6.2882 | -84.0022
6.2885 | -84.0013
6.2888 | -84.0004
6.2892 | -84.0016
6.2895 | -84.002
6.2898 | -84.002
6.2902 | -84.0018
6.2905 | -84.0019
6.2908 | -84.0015
6.2911 | -84.0009
6.2915 | -84.0005
6.2918 | -84.001
6.2921 | -84.0009
6.2925 | -84.0007
6.2928 | -84.0004
6.2931 | -83.9998
6.2934 | -83.9994
6.2938 | -83.9995
6.2941 | -83.9997
6.2944 | -83.9994
6.2948 | -83.9996
6.2951 | -83.9989
6.2954 | -83.9988
6.2957 | -83.9992
6.2961 | -83.9989
6.2964 | -83.9986
6.2967 | -83.9987
6.2971 | -83.9991
6.2974 | -83.9998
6.2977 | -84.0004
6.2981 | -84.0001
6.2984 | -83.9999
6.2987 | -83.9998
6.299 | -83.9995
6.2994 | -83.9992
6.2997 | -83.9989
6.3 | -83.9979
6.3004 | -83.9978
6.3007 | -83.9972
6.301 | -83.9947
6.3013 | -83.9931
6.3017 | -83.9928
6.302 | -83.9929
6.3023 | -83.9927
6.3027 | -83.9926
6.303 | -83.9924
6.3033 | -83.9923
6.3036 | -83.9919
6.304 | -83.9918
6.3043 | -83.9923
6.3046 | -83.9924
6.305 | -83.9924
6.3053 | -83.9927
6.3056 | -83.9923
6.306 | -83.9916
6.3063 | -83.9915
6.3066 | -83.991
6.3069 | -83.9908
6.3073 | -83.9915
6.3076 | -83.991
6.3079 | -83.9906
6.3083 | -83.9901
6.3086 | -83.9895
6.3089 | -83.9892
6.3092 | -83.9904
6.3096 | -83.9912
6.3099 | -83.9903
6.3102 | -83.99
6.3106 | -83.99
6.3109 | -83.9896
6.3112 | -83.9892
6.3115 | -83.9894
6.3119 | -83.9897
6.3122 | -83.989
6.3125 | -83.9882
6.3129 | -83.9879
6.3132 | -83.9874
6.3135 | -83.9876
6.3139 | -83.9876
6.3142 | -83.9872
6.3145 | -83.9867
6.3148 | -83.9861
6.3152 | -83.9849
6.3155 | -83.9846
6.3158 | -83.9856
6.3162 | -83.9848
6.3165 | -83.9837
6.3168 | -83.983
6.3171 | -83.9821
6.3175 | -83.9803
6.3178 | -83.9796
6.3181 | -83.9803
6.3185 | -83.9798
6.3188 | -83.9774
6.3191 | -83.9762
6.3194 | -83.9752
6.3198 | -83.9745
6.3201 | -83.9745
6.3204 | -83.9748
6.3208 | -83.9754
6.3211 | -83.9749
6.3214 | -83.9754
6.3218 | -83.9748
6.3221 | -83.9741
6.3224 | -83.9734
6.3227 | -83.9725
6.3231 | -83.972
6.3234 | -83.9713
6.3237 | -83.9709
6.3241 | -83.9716
6.3244 | -83.9714
6.3247 | -83.9722
6.325 | -83.9716
6.3254 | -83.9708
6.3257 | -83.9713
6.326 | -83.9702
6.3264 | -83.9694
6.3267 | -83.9691
6.327 | -83.9687
6.3273 | -83.9681
6.3277 | -83.9682
6.328 | -83.9687
6.3283 | -83.9684
6.3287 | -83.9672
6.329 | -83.9662
6.3293 | -83.9655
6.3297 | -83.9649
6.33 | -83.9646
6.3303 | -83.9648
6.3306 | -83.9647
6.331 | -83.9638
6.3313 | -83.9628
6.3316 | -83.9621
6.332 | -83.9613
6.3323 | -83.9609
6.3326 | -83.961
6.3329 | -83.9604
6.3333 | -83.9601
6.3336 | -83.9597
6.3339 | -83.9591
6.3343 | -83.9584
6.3346 | -83.9582
6.3349 | -83.959
6.3352 | -83.9594
6.3356 | -83.9588
6.3359 | -83.9588
6.3362 | -83.9586
6.3366 | -83.9583
6.3369 | -83.9578
6.3372 | -83.9578
6.3376 | -83.9578
6.3379 | -83.9571
6.3382 | -83.9565
6.3385 | -83.9561
6.3389 | -83.9555
6.3392 | -83.9552
6.3395 | -83.9554
6.3399 | -83.9549
6.3402 | -83.9542
6.3405 | -83.9534
6.3408 | -83.9532
6.3412 | -83.9525
6.3415 | -83.9523
6.3418 | -83.9522
6.3422 | -83.9519
6.3425 | -83.9511
6.3428 | -83.9504
6.3431 | -83.9489
6.3435 | -83.9484
6.3438 | -83.9475
6.3441 | -83.9468
6.3445 | -83.9472
6.3448 | -83.9467
6.3451 | -83.9457
6.3454 | -83.9444
6.3458 | -83.9432
6.3461 | -83.943
6.3464 | -83.9424
6.3468 | -83.9421
6.3471 | -83.9419
6.3474 | -83.9414
6.3478 | -83.9407
6.3481 | -83.9399
6.3484 | -83.9391
6.3487 | -83.9377
6.3491 | -83.9366
6.3494 | -83.9354
6.3497 | -83.9345
6.3501 | -83.9337
6.3504 | -83.9331
6.3507 | -83.9322
6.351 | -83.9309
6.3514 | -83.9299
6.3517 | -83.9292
6.352 | -83.9287
6.3524 | -83.9297
6.3527 | -83.93
6.353 | -83.9301
6.3533 | -83.9295
6.3537 | -83.9292
6.354 | -83.9297
6.3543 | -83.9293
6.3547 | -83.9283
6.355 | -83.9269
6.3553 | -83.9253
6.3557 | -83.9244
6.356 | -83.9232
6.3563 | -83.9224
6.3566 | -83.9231
6.357 | -83.9232
6.3573 | -83.9218
6.3576 | -83.9202
6.358 | -83.9192
6.3583 | -83.9182
6.3586 | -83.9173
6.3589 | -83.9167
6.3593 | -83.9154
6.3596 | -83.9143
6.3599 | -83.9142
6.3603 | -83.914
6.3606 | -83.9132
6.3609 | -83.9138
6.3612 | -83.9135
6.3616 | -83.9126
6.3619 | -83.9122
6.3622 | -83.9115
6.3626 | -83.9114
6.3629 | -83.911
6.3632 | -83.9104
6.3636 | -83.9102
6.3639 | -83.9095
6.3642 | -83.9098
6.3645 | -83.9108
6.3649 | -83.9104
6.3652 | -83.9101
6.3655 | -83.9101
6.3659 | -83.9096
6.3662 | -83.9091
6.3665 | -83.9101
6.3668 | -83.9106
6.3672 | -83.9096
6.3675 | -83.9084
6.3678 | -83.9084
6.3682 | -83.9083
6.3685 | -83.9076
6.3688 | -83.908
6.3691 | -83.9087
6.3695 | -83.9081
6.3698 | -83.9075
6.3701 | -83.9083
6.3705 | -83.9089
6.3708 | -83.9083
6.3711 | -83.9078
6.3715 | -83.9073
6.3718 | -83.9076
6.3721 | -83.9071
6.3724 | -83.9068
6.3728 | -83.9061
6.3731 | -83.9053
6.3734 | -83.9045
6.3738 | -83.904
6.3741 | -83.9034
6.3744 | -83.9027
6.3747 | -83.9013
6.3751 | -83.8997
6.3754 | -83.8979
6.3757 | -83.8963
6.3761 | -83.895
6.3764 | -83.8942
6.3767 | -83.8937
6.377 | -83.8937
6.3774 | -83.8935
6.3777 | -83.893
6.378 | -83.8924
6.3784 | -83.8918
6.3787 | -83.8914
6.379 | -83.8911
6.3794 | -83.8908
6.3797 | -83.8901
6.38 | -83.8892
6.3803 | -83.8888
6.3807 | -83.8883
6.381 | -83.8884
6.3813 | -83.8887
6.3817 | -83.8887
6.382 | -83.8884
6.3823 | -83.8872
6.3826 | -83.8868
6.383 | -83.8864
6.3833 | -83.8859
6.3836 | -83.8853
6.384 | -83.884
6.3843 | -83.8835
6.3846 | -83.8827
6.3849 | -83.8822
6.3853 | -83.8831
6.3856 | -83.8836
6.3859 | -83.8838
6.3863 | -83.8833
6.3866 | -83.8822
6.3869 | -83.8813
6.3873 | -83.8807
6.3876 | -83.8803
6.3879 | -83.8795
6.3882 | -83.8794
6.3886 | -83.8793
6.3889 | -83.8794
6.3892 | -83.8801
6.3896 | -83.8801
6.3899 | -83.8811
6.3902 | -83.882
6.3905 | -83.8823
6.3909 | -83.8827
6.3912 | -83.8827
6.3915 | -83.8831
6.3919 | -83.8837
6.3922 | -83.8835
6.3925 | -83.8828
6.3928 | -83.8826
6.3932 | -83.8816
6.3935 | -83.8793
6.3938 | -83.879
6.3942 | -83.8786
6.3945 | -83.8786
6.3948 | -83.8782
6.3952 | -83.8779
6.3955 | -83.8776
6.3958 | -83.8775
6.3961 | -83.8776
6.3965 | -83.8768
6.3968 | -83.8765
6.3971 | -83.8747
6.3975 | -83.8734
6.3978 | -83.8729
6.3981 | -83.8723
6.3984 | -83.8715
6.3988 | -83.8708
6.3991 | -83.8698
6.3994 | -83.8692
6.3998 | -83.8687
6.4001 | -83.8683
6.4004 | -83.868
6.4007 | -83.8672
6.4011 | -83.8673
6.4014 | -83.8666
6.4017 | -83.866
6.4021 | -83.8667
6.4024 | -83.8668
6.4027 | -83.8663
6.403 | -83.8653
6.4034 | -83.8646
6.4037 | -83.8642
6.404 | -83.864
6.4044 | -83.8634
6.4047 | -83.863
6.405 | -83.8627
6.4054 | -83.8626
6.4057 | -83.8626
6.406 | -83.8626
6.4063 | -83.8622
6.4067 | -83.8619
6.407 | -83.8618
6.4073 | -83.8614
6.4077 | -83.8608
6.408 | -83.8607
6.4083 | -83.8608
6.4086 | -83.8603
6.409 | -83.8598
6.4093 | -83.8598
6.4096 | -83.8587
6.41 | -83.857
6.4103 | -83.8566
6.4106 | -83.856
6.4109 | -83.8556
6.4113 | -83.8552
6.4116 | -83.8548
6.4119 | -83.8549
6.4123 | -83.8546
6.4126 | -83.8539
6.4129 | -83.8533
6.4133 | -83.8527
6.4136 | -83.8525
6.4139 | -83.8522
6.4142 | -83.852
6.4146 | -83.852
6.4149 | -83.8519
6.4152 | -83.8522
6.4156 | -83.852
6.4159 | -83.8519
6.4162 | -83.8519
6.4165 | -83.8517
6.4169 | -83.852
6.4172 | -83.8517
6.4175 | -83.8517
6.4179 | -83.8512
6.4182 | -83.8497
6.4185 | -83.8487
6.4188 | -83.8487
6.4192 | -83.8478
6.4195 | -83.8481
6.4198 | -83.8485
6.4202 | -83.8484
6.4205 | -83.8478
6.4208 | -83.8475
6.4212 | -83.8477
6.4215 | -83.8467
6.4218 | -83.8462
6.4221 | -83.8449
6.4225 | -83.8451
6.4228 | -83.8465
6.4231 | -83.8485
6.4235 | -83.8489
6.4238 | -83.8495
6.4241 | -83.8497
6.4244 | -83.8495
6.4248 | -83.8489
6.4251 | -83.8479
6.4254 | -83.8471
6.4258 | -83.8466
6.4261 | -83.8462
6.4264 | -83.8457
6.4267 | -83.8454
6.4271 | -83.8452
6.4274 | -83.8446
6.4277 | -83.8446
6.4281 | -83.8446
6.4284 | -83.8443
6.4287 | -83.8442
6.4291 | -83.8452
6.4294 | -83.8449
6.4297 | -83.8446
6.43 | -83.8447
6.4304 | -83.8452
6.4307 | -83.8455
6.431 | -83.8455
6.4314 | -83.8456
6.4317 | -83.8459
6.432 | -83.8462
6.4323 | -83.8468
6.4327 | -83.8466
6.433 | -83.8464
6.4333 | -83.8465
6.4337 | -83.8465
6.434 | -83.8465
6.4343 | -83.847
6.4346 | -83.8472
6.435 | -83.8479
6.4353 | -83.8478
6.4356 | -83.8484
6.436 | -83.8492
6.4363 | -83.849
6.4366 | -83.849
6.437 | -83.849
6.4373 | -83.8488
6.4376 | -83.8484
6.4379 | -83.848
6.4383 | -83.8488
6.4386 | -83.8494
6.4389 | -83.8487
6.4393 | -83.8484
6.4396 | -83.8487
6.4399 | -83.8489
6.4402 | -83.849
6.4406 | -83.849
6.4409 | -83.8493
6.4412 | -83.8495
6.4416 | -83.8497
6.4419 | -83.8497
6.4422 | -83.8498
6.4425 | -83.8497
6.4429 | -83.8497
6.4432 | -83.8504
6.4435 | -83.8516
6.4439 | -83.8521
6.4442 | -83.8523
6.4445 | -83.8524
6.4448 | -83.8528
6.4452 | -83.853
6.4455 | -83.8538
6.4458 | -83.854
6.4462 | -83.8544
6.4465 | -83.8544
6.4468 | -83.8543
6.4472 | -83.8537
6.4475 | -83.8533
6.4478 | -83.854
6.4481 | -83.8537
6.4485 | -83.8532
6.4488 | -83.8532
6.4491 | -83.8533
6.4495 | -83.8539
6.4498 | -83.8546
6.4501 | -83.8548
6.4504 | -83.8555
6.4508 | -83.8555
6.4511 | -83.856
6.4514 | -83.8564
6.4518 | -83.8564
6.4521 | -83.8567
6.4524 | -83.857
6.4527 | -83.8565
6.4531 | -83.8563
6.4534 | -83.8566
6.4537 | -83.8568
6.4541 | -83.8572
6.4544 | -83.8568
6.4547 | -83.8558
6.4551 | -83.8558
6.4554 | -83.8565
6.4557 | -83.8572
6.456 | -83.8576
6.4564 | -83.8584
6.4567 | -83.8589
6.457 | -83.8591
6.4574 | -83.8604
6.4577 | -83.8612
6.458 | -83.8627
6.4583 | -83.8635
6.4587 | -83.8642
6.459 | -83.8645
6.4593 | -83.865
6.4597 | -83.8654
6.46 | -83.8659
6.4603 | -83.8663
6.4606 | -83.867
6.461 | -83.8678
6.4613 | -83.868
6.4616 | -83.8684
6.462 | -83.8688
6.4623 | -83.8696
6.4626 | -83.8703
6.463 | -83.8709
6.4633 | -83.8724
6.4636 | -83.8733
6.4639 | -83.8737
6.4643 | -83.874
6.4646 | -83.8737
6.4649 | -83.8745
6.4653 | -83.876
6.4656 | -83.8767
6.4659 | -83.8773
6.4662 | -83.878
6.4666 | -83.8787
6.4669 | -83.8796
6.4672 | -83.8802
6.4676 | -83.8807
6.4679 | -83.8808
6.4682 | -83.881
6.4685 | -83.8816
6.4689 | -83.8813
6.4692 | -83.8817
6.4695 | -83.8821
6.4699 | -83.8827
6.4702 | -83.8832
6.4705 | -83.8842
6.4709 | -83.8867
6.4712 | -83.8891
6.4715 | -83.8898
6.4718 | -83.8895
6.4722 | -83.8913
6.4725 | -83.8921
6.4728 | -83.8928
6.4732 | -83.8932
6.4735 | -83.8936
6.4738 | -83.8944
6.4741 | -83.8947
6.4745 | -83.895
6.4748 | -83.8956
6.4751 | -83.8955
6.4755 | -83.8962
6.4758 | -83.8971
6.4761 | -83.8979
6.4764 | -83.8993
6.4768 | -83.9002
6.4771 | -83.9008
6.4774 | -83.9021
6.4778 | -83.9037
6.4781 | -83.9049
6.4784 | -83.9055
6.4788 | -83.906
6.4791 | -83.9065
6.4794 | -83.9074
6.4797 | -83.9079
6.4801 | -83.9078
6.4804 | -83.908
6.4807 | -83.9085
6.4811 | -83.9084
6.4814 | -83.9086
6.4817 | -83.9093
6.482 | -83.911
6.4824 | -83.9136
6.4827 | -83.9153
6.483 | -83.9165
6.4834 | -83.9174
6.4837 | -83.9183
6.484 | -83.919
6.4843 | -83.9198
6.4847 | -83.921
6.485 | -83.9211
6.4853 | -83.9216
6.4857 | -83.9218
6.486 | -83.9228
6.4863 | -83.9233
6.4867 | -83.924
6.487 | -83.9246
6.4873 | -83.9253
6.4876 | -83.9258
6.488 | -83.9264
6.4883 | -83.927
6.4886 | -83.9278
6.489 | -83.9285
6.4893 | -83.9295
6.4896 | -83.9303
6.4899 | -83.9311
6.4903 | -83.9326
6.4906 | -83.9337
6.4909 | -83.9336
6.4913 | -83.9339
6.4916 | -83.9336
6.4919 | -83.9338
6.4922 | -83.9348
6.4926 | -83.9356
6.4929 | -83.9362
6.4932 | -83.937
6.4936 | -83.937
6.4939 | -83.9379
6.4942 | -83.9393
6.4945 | -83.94
6.4949 | -83.9404
6.4952 | -83.941
6.4955 | -83.9427
6.4959 | -83.9436
6.4962 | -83.9441
6.4965 | -83.9451
6.4969 | -83.9461
6.4972 | -83.9473
6.4975 | -83.9487
6.4978 | -83.9496
6.4982 | -83.9504
6.4985 | -83.9512
6.4988 | -83.952
6.4992 | -83.9528
6.4995 | -83.9542
6.4998 | -83.9553
6.5001 | -83.9557
6.5005 | -83.9563
6.5008 | -83.9572
6.5011 | -83.9578
6.5015 | -83.9587
6.5018 | -83.9594
6.5021 | -83.9603
6.5024 | -83.961
6.5028 | -83.9617
6.5031 | -83.9622
6.5034 | -83.9635
6.5038 | -83.9645
6.5041 | -83.9647
6.5044 | -83.9646
6.5048 | -83.9651
6.5051 | -83.9641
6.5054 | -83.9649
6.5057 | -83.9677
6.5061 | -83.9701
6.5064 | -83.9711
6.5067 | -83.9721
6.5071 | -83.973
6.5074 | -83.9739
6.5077 | -83.9749
6.508 | -83.9752
6.5084 | -83.9762
6.5087 | -83.9769
6.509 | -83.9778
6.5094 | -83.9788
6.5097 | -83.9796
6.51 | -83.9807
6.5103 | -83.9817
6.5107 | -83.9824
6.511 | -83.9833
6.5113 | -83.9843
6.5117 | -83.9854
6.512 | -83.9874
6.5123 | -83.9895
6.5127 | -83.9908
6.513 | -83.9912
6.5133 | -83.9916
6.5136 | -83.9923
6.514 | -83.9931
6.5143 | -83.994
6.5146 | -83.995
6.515 | -83.9964
6.5153 | -83.9979
6.5156 | -83.9992
6.5159 | -84.0006
6.5163 | -84.0019
6.5166 | -84.003
6.5169 | -84.0038
6.5173 | -84.0037
6.5176 | -84.0036
6.5179 | -84.0055
6.5182 | -84.0074
6.5186 | -84.0085
6.5189 | -84.0094
6.5192 | -84.0103
6.5196 | -84.0104
6.5199 | -84.0108
6.5202 | -84.0116
6.5206 | -84.0131
6.5209 | -84.0153
6.5212 | -84.0174
6.5215 | -84.0187
6.5219 | -84.0195
6.5222 | -84.0205
6.5225 | -84.0214
6.5229 | -84.0225
6.5232 | -84.0237
6.5235 | -84.0248
6.5238 | -84.026
6.5242 | -84.0268
6.5245 | -84.0279
6.5248 | -84.0281
6.5252 | -84.0292
6.5255 | -84.0296
6.5258 | -84.03
6.5261 | -84.031
6.5265 | -84.0316
6.5268 | -84.0325
6.5271 | -84.0339
6.5275 | -84.0352
6.5278 | -84.0362
6.5281 | -84.0371
6.5285 | -84.0381
6.5288 | -84.0391
6.5291 | -84.0404
6.5294 | -84.0413
6.5298 | -84.0421
6.5301 | -84.0432
6.5304 | -84.045
6.5308 | -84.0461
6.5311 | -84.0466
6.5314 | -84.0474
6.5317 | -84.0483
6.5321 | -84.0489
6.5324 | -84.0495
6.5327 | -84.0509
6.5331 | -84.0522
6.5334 | -84.0538
6.5337 | -84.0552
6.534 | -84.0566
6.5344 | -84.0578
6.5347 | -84.0593
6.535 | -84.0603
6.5354 | -84.0606
6.5357 | -84.0609
6.536 | -84.0624
6.5363 | -84.0637
6.5367 | -84.0649
6.537 | -84.0658
6.5373 | -84.0667
6.5377 | -84.0678
6.538 | -84.0685
6.5383 | -84.0695
6.5387 | -84.0706
6.539 | -84.0727
6.5393 | -84.0744
6.5396 | -84.0754
6.54 | -84.0771
6.5403 | -84.0793
6.5406 | -84.0805
6.541 | -84.082
6.5413 | -84.0827
6.5416 | -84.0847
6.5419 | -84.0862
6.5423 | -84.0871
6.5426 | -84.0881
6.5429 | -84.0899
6.5433 | -84.0914
6.5436 | -84.0937
6.5439 | -84.0949
6.5442 | -84.096
6.5446 | -84.0973
6.5449 | -84.0988
6.5452 | -84.0995
6.5456 | -84.1003
6.5459 | -84.1012
6.5462 | -84.1011
6.5466 | -84.1016
6.5469 | -84.1025
6.5472 | -84.1034
6.5475 | -84.1044
6.5479 | -84.1054
6.5482 | -84.1062
6.5485 | -84.1064
6.5489 | -84.1067
6.5492 | -84.1073
6.5495 | -84.1082
6.5498 | -84.1094
6.5502 | -84.11
6.5505 | -84.1107
6.5508 | -84.1111
6.5512 | -84.1119
6.5515 | -84.1129
6.5518 | -84.1144
6.5521 | -84.1163
6.5525 | -84.1178
6.5528 | -84.119
6.5531 | -84.1202
6.5535 | -84.1214
6.5538 | -84.1222
6.5541 | -84.1234
6.5545 | -84.1246
6.5548 | -84.1265
6.5551 | -84.1281
6.5554 | -84.129
6.5558 | -84.1304
6.5561 | -84.1322
6.5564 | -84.1335
6.5568 | -84.1346
6.5571 | -84.135
6.5574 | -84.1355
6.5577 | -84.136
6.5581 | -84.1371
6.5584 | -84.1377
6.5587 | -84.139
6.5591 | -84.1393
6.5594 | -84.1405
6.5597 | -84.1426
6.56 | -84.1457
6.5604 | -84.1472
6.5607 | -84.1477
6.561 | -84.1484
6.5614 | -84.1488
6.5617 | -84.1496
6.562 | -84.1509
6.5624 | -84.1518
6.5627 | -84.153
6.563 | -84.1542
6.5633 | -84.1555
6.5637 | -84.1569
6.564 | -84.1579
6.5643 | -84.1587
6.5647 | -84.1595
6.565 | -84.1603
6.5653 | -84.1614
6.5656 | -84.1623
6.566 | -84.1627
6.5663 | -84.1639
6.5666 | -84.165
6.567 | -84.1658
6.5673 | -84.1667
6.5676 | -84.1677
6.5679 | -84.1686
6.5683 | -84.1697
6.5686 | -84.1709
6.5689 | -84.1722
6.5693 | -84.1734
6.5696 | -84.1746
6.5699 | -84.1754
6.5703 | -84.1756
6.5706 | -84.1769
6.5709 | -84.1782
6.5712 | -84.1801
6.5716 | -84.1814
6.5719 | -84.1828
6.5722 | -84.1844
6.5726 | -84.1854
6.5729 | -84.1866
6.5732 | -84.1876
6.5735 | -84.1887
6.5739 | -84.1902
6.5742 | -84.1913
6.5745 | -84.1925
6.5749 | -84.1935
6.5752 | -84.1946
6.5755 | -84.1959
6.5758 | -84.1976
6.5762 | -84.1992
6.5765 | -84.2009
6.5768 | -84.2025
6.5772 | -84.2036
6.5775 | -84.2049
6.5778 | -84.2065
6.5782 | -84.208
6.5785 | -84.2092
6.5788 | -84.2102
6.5791 | -84.2108
6.5795 | -84.2111
6.5798 | -84.2119
6.5801 | -84.2131
6.5805 | -84.2149
6.5808 | -84.2155
6.5811 | -84.2172
6.5814 | -84.219
6.5818 | -84.2204
6.5821 | -84.2218
6.5824 | -84.2232
6.5828 | -84.2243
6.5831 | -84.2257
6.5834 | -84.2274
6.5837 | -84.2296
6.5841 | -84.2309
6.5844 | -84.2314
6.5847 | -84.2323
6.5851 | -84.2336
6.5854 | -84.2357
6.5857 | -84.2374
6.5861 | -84.2386
6.5864 | -84.24
6.5867 | -84.2419
6.587 | -84.2434
6.5874 | -84.2446
6.5877 | -84.2458
6.588 | -84.247
6.5884 | -84.2483
6.5887 | -84.2494
6.589 | -84.2503
6.5893 | -84.2517
6.5897 | -84.2527
6.59 | -84.2537
6.5903 | -84.2549
6.5907 | -84.2558
6.591 | -84.257
6.5913 | -84.2581
6.5916 | -84.2593
6.592 | -84.2608
6.5923 | -84.262
6.5926 | -84.263
6.593 | -84.2643
6.5933 | -84.2656
6.5936 | -84.2682
6.5939 | -84.2705
6.5943 | -84.2713
6.5946 | -84.2726
6.5949 | -84.2739
6.5953 | -84.2752
6.5956 | -84.277
6.5959 | -84.2782
6.5963 | -84.2794
6.5966 | -84.2813
6.5969 | -84.283
6.5972 | -84.2844
6.5976 | -84.2858
6.5979 | -84.2877
6.5982 | -84.2898
6.5986 | -84.2919
6.5989 | -84.2933
6.5992 | -84.2945
6.5995 | -84.2962
6.5999 | -84.298
6.6002 | -84.3
6.6005 | -84.3014
6.6009 | -84.3032
6.6012 | -84.3058
6.6015 | -84.3076
6.6018 | -84.3089
6.6022 | -84.311
6.6025 | -84.3132
6.6028 | -84.3152
6.6032 | -84.317
6.6035 | -84.3187
6.6038 | -84.3194
6.6042 | -84.3208
6.6045 | -84.3225
6.6048 | -84.3239
6.6051 | -84.3256
6.6055 | -84.3274
6.6058 | -84.3292
6.6061 | -84.3306
6.6065 | -84.3325
6.6068 | -84.334
6.6071 | -84.3349
6.6074 | -84.3365
6.6078 | -84.3386
6.6081 | -84.3397
6.6084 | -84.3408
6.6088 | -84.3421
6.6091 | -84.3434
6.6094 | -84.3449
6.6097 | -84.3465
6.6101 | -84.3486
6.6104 | -84.35
6.6107 | -84.3512
6.6111 | -84.3526
6.6114 | -84.3539
6.6117 | -84.355
6.6121 | -84.3563
6.6124 | -84.3584
6.6127 | -84.3602
6.613 | -84.3607
6.6134 | -84.3616
6.6137 | -84.363
6.614 | -84.3649
6.6144 | -84.3653
6.6147 | -84.3665
6.615 | -84.3674
6.6153 | -84.3687
6.6157 | -84.3706
6.616 | -84.3723
6.6163 | -84.3743
6.6167 | -84.3757
6.617 | -84.3771
6.6173 | -84.3787
6.6176 | -84.3812
6.618 | -84.3838
6.6183 | -84.3856
6.6186 | -84.3867
6.619 | -84.388
6.6193 | -84.3898
6.6196 | -84.3927
6.62 | -84.3934
6.6203 | -84.3943
6.6206 | -84.3957
6.6209 | -84.397
6.6213 | -84.399
6.6216 | -84.4011
6.6219 | -84.4023
6.6223 | -84.4037
6.6226 | -84.4052
6.6229 | -84.4063
6.6232 | -84.4077
6.6236 | -84.409
6.6239 | -84.4101
6.6242 | -84.4114
6.6246 | -84.4126
6.6249 | -84.4143
6.6252 | -84.4165
6.6255 | -84.418
6.6259 | -84.4193
6.6262 | -84.4208
6.6265 | -84.4229
6.6269 | -84.4257
6.6272 | -84.4267
6.6275 | -84.4275
6.6278 | -84.4303
6.6282 | -84.4318
6.6285 | -84.4338
6.6288 | -84.4351
6.6292 | -84.4373
6.6295 | -84.4387
6.6298 | -84.44
6.6302 | -84.4412
6.6305 | -84.4426
6.6308 | -84.4443
6.6311 | -84.446
6.6315 | -84.4469
6.6318 | -84.4487
6.6321 | -84.4503
6.6325 | -84.4522
6.6328 | -84.454
6.6331 | -84.4554
6.6334 | -84.4568
6.6338 | -84.4599
6.6341 | -84.4638
6.6344 | -84.4657
6.6348 | -84.4674
6.6351 | -84.4689
6.6354 | -84.4699
6.6357 | -84.4709
6.6361 | -84.4724
6.6364 | -84.4738
6.6367 | -84.4755
6.6371 | -84.4786
6.6374 | -84.4814
6.6377 | -84.483
6.6381 | -84.4844
6.6384 | -84.4861
6.6387 | -84.4875
6.639 | -84.4884
6.6394 | -84.4904
6.6397 | -84.4923
6.64 | -84.494
6.6404 | -84.4956
6.6407 | -84.4967
6.641 | -84.4967
6.6413 | -84.4979
6.6417 | -84.4995
6.642 | -84.5006
6.6423 | -84.5022
6.6427 | -84.5049
6.643 | -84.5064
6.6433 | -84.5084
6.6436 | -84.5109
6.644 | -84.5126
6.6443 | -84.5141
6.6446 | -84.5154
6.645 | -84.5171
6.6453 | -84.519
6.6456 | -84.5207
6.646 | -84.5227
6.6463 | -84.5242
6.6466 | -84.5248
6.6469 | -84.5258
6.6473 | -84.5279
6.6476 | -84.529
6.6479 | -84.5305
6.6483 | -84.5318
6.6486 | -84.5335
6.6489 | -84.5345
6.6492 | -84.5352
6.6496 | -84.5377
6.6499 | -84.5397
6.6502 | -84.5399
6.6506 | -84.5412
6.6509 | -84.5437
6.6512 | -84.5453
6.6515 | -84.5467
6.6519 | -84.5484
6.6522 | -84.5498
6.6525 | -84.5519
6.6529 | -84.5538
6.6532 | -84.5549
6.6535 | -84.5564
6.6539 | -84.558
6.6542 | -84.5598
6.6545 | -84.5607
6.6548 | -84.5628
6.6552 | -84.5642
6.6555 | -84.5657
6.6558 | -84.5674
6.6562 | -84.569
6.6565 | -84.5708
6.6568 | -84.5727
6.6571 | -84.5747
6.6575 | -84.5765
6.6578 | -84.5782
6.6581 | -84.5795
6.6585 | -84.5804
6.6588 | -84.5815
6.6591 | -84.5832
6.6594 | -84.5851
6.6598 | -84.5864
6.6601 | -84.5877
6.6604 | -84.5891
6.6608 | -84.5907
6.6611 | -84.5923
6.6614 | -84.5936
6.6618 | -84.595
6.6621 | -84.5963
6.6624 | -84.5978
6.6627 | -84.5986
6.6631 | -84.6009
6.6634 | -84.6026
6.6637 | -84.6054
6.6641 | -84.6091
6.6644 | -84.6108
6.6647 | -84.6121
6.665 | -84.6134
6.6654 | -84.6147
6.6657 | -84.6159
6.666 | -84.6172
6.6664 | -84.6185
6.6667 | -84.6197
6.667 | -84.6212
6.6673 | -84.6228
6.6677 | -84.6242
6.668 | -84.6256
6.6683 | -84.627
6.6687 | -84.6288
6.669 | -84.6313
6.6693 | -84.6331
6.6697 | -84.6349
6.67 | -84.6367
6.6703 | -84.6381
6.6706 | -84.6395
6.671 | -84.642
6.6713 | -84.6442
6.6716 | -84.6456
6.672 | -84.6471
6.6723 | -84.6489
6.6726 | -84.6508
6.6729 | -84.6531
6.6733 | -84.6538
6.6736 | -84.6551
6.6739 | -84.6566
6.6743 | -84.6579
6.6746 | -84.6591
6.6749 | -84.6606
6.6752 | -84.6624
6.6756 | -84.6639
6.6759 | -84.6657
6.6762 | -84.6671
6.6766 | -84.6686
6.6769 | -84.6699
6.6772 | -84.6714
6.6776 | -84.6729
6.6779 | -84.6738
6.6782 | -84.6757
6.6785 | -84.6772
6.6789 | -84.6785
6.6792 | -84.6806
6.6795 | -84.6848
6.6799 | -84.6874
6.6802 | -84.689
6.6805 | -84.6907
6.6808 | -84.6922
6.6812 | -84.6938
6.6815 | -84.6959
6.6818 | -84.6978
6.6822 | -84.6993
6.6825 | -84.7011
6.6828 | -84.7028
6.6831 | -84.7045
6.6835 | -84.7062
6.6838 | -84.7076
6.6841 | -84.709
6.6845 | -84.7115
6.6848 | -84.7141
6.6851 | -84.7172
6.6854 | -84.7193
6.6858 | -84.7219
6.6861 | -84.7238
6.6864 | -84.7254
6.6868 | -84.7262
6.6871 | -84.728
6.6874 | -84.7294
6.6878 | -84.7311
6.6881 | -84.7329
6.6884 | -84.7345
6.6887 | -84.7361
6.6891 | -84.7379
6.6894 | -84.7396
6.6897 | -84.741
6.6901 | -84.7421
6.6904 | -84.7435
6.6907 | -84.7453
6.691 | -84.7469
6.6914 | -84.7491
6.6917 | -84.7509
6.692 | -84.7527
6.6924 | -84.7549
6.6927 | -84.7571
6.693 | -84.7587
6.6933 | -84.7601
6.6937 | -84.7616
6.694 | -84.7633
6.6943 | -84.7645
6.6947 | -84.7666
6.695 | -84.7677
6.6953 | -84.7687
6.6957 | -84.7707
6.696 | -84.7726
6.6963 | -84.7737
6.6966 | -84.7748
6.697 | -84.7759
6.6973 | -84.7774
6.6976 | -84.7787
6.698 | -84.7801
6.6983 | -84.7813
6.6986 | -84.7832
6.6989 | -84.7848
6.6993 | -84.7864
6.6996 | -84.7881
6.6999 | -84.7904
6.7003 | -84.7921
6.7006 | -84.7938
6.7009 | -84.7964
6.7012 | -84.7987
6.7016 | -84.7992
6.7019 | -84.8
6.7022 | -84.8016
6.7026 | -84.8034
6.7029 | -84.805
6.7032 | -84.8066
6.7036 | -84.8085
6.7039 | -84.8107
6.7042 | -84.812
6.7045 | -84.8141
6.7049 | -84.816
6.7052 | -84.8172
6.7055 | -84.819
6.7059 | -84.821
6.7062 | -84.8227
6.7065 | -84.8239
6.7068 | -84.8254
6.7072 | -84.8271
6.7075 | -84.8288
6.7078 | -84.8302
6.7082 | -84.8321
6.7085 | -84.8348
6.7088 | -84.8363
6.7091 | -84.8379
6.7095 | -84.8395
6.7098 | -84.8414
6.7101 | -84.8434
6.7105 | -84.8454
6.7108 | -84.8466
6.7111 | -84.8467
6.7115 | -84.8471
6.7118 | -84.8477
6.7121 | -84.8489
6.7124 | -84.8504
6.7128 | -84.8522
6.7131 | -84.8534
6.7134 | -84.8555
6.7138 | -84.8576
6.7141 | -84.8598
6.7144 | -84.8616
6.7147 | -84.8636
6.7151 | -84.8667
6.7154 | -84.8684
6.7157 | -84.8702
6.7161 | -84.873
6.7164 | -84.876
6.7167 | -84.8791
6.717 | -84.8813
6.7174 | -84.8829
6.7177 | -84.8844
6.718 | -84.8865
6.7184 | -84.8895
6.7187 | -84.8911
6.719 | -84.8916
6.7194 | -84.8945
6.7197 | -84.8972
6.72 | -84.8989
6.7203 | -84.9007
6.7207 | -84.9022
6.721 | -84.9037
6.7213 | -84.9058
6.7217 | -84.9078
6.722 | -84.9095
6.7223 | -84.9111
6.7226 | -84.9124
6.723 | -84.9136
6.7233 | -84.9149
6.7236 | -84.9163
6.724 | -84.9179
6.7243 | -84.92
6.7246 | -84.9222
6.7249 | -84.9236
6.7253 | -84.9252
6.7256 | -84.9267
6.7259 | -84.9279
6.7263 | -84.9296
6.7266 | -84.9313
6.7269 | -84.9318
6.7272 | -84.9324
6.7276 | -84.9336
6.7279 | -84.9354
6.7282 | -84.9369
6.7286 | -84.9387
6.7289 | -84.9411
6.7292 | -84.9424
6.7296 | -84.943
6.7299 | -84.9441
6.7302 | -84.9455
6.7305 | -84.9468
6.7309 | -84.9485
6.7312 | -84.95
6.7315 | -84.9514
6.7319 | -84.9531
6.7322 | -84.955
6.7325 | -84.9561
6.7328 | -84.958
6.7332 | -84.9593
6.7335 | -84.9613
6.7338 | -84.9633
6.7342 | -84.9647
6.7345 | -84.9658
6.7348 | -84.9673
6.7351 | -84.9693
6.7355 | -84.9716
6.7358 | -84.9726
6.7361 | -84.9737
6.7365 | -84.9751
6.7368 | -84.9765
6.7371 | -84.9782
6.7375 | -84.9797
6.7378 | -84.9815
6.7381 | -84.9828
6.7384 | -84.9838
6.7388 | -84.9852
6.7391 | -84.987
6.7394 | -84.9886
6.7398 | -84.9898
6.7401 | -84.991
6.7404 | -84.9926
6.7407 | -84.9946
6.7411 | -84.9971
6.7414 | -84.9993
6.7417 | -85.0006
6.7421 | -85.0019
6.7424 | -85.001
6.7427 | -85.0015
6.743 | -85.0021
6.7434 | -85.0037
6.7437 | -85.0058
6.744 | -85.0077
6.7444 | -85.0095
6.7447 | -85.0108
6.745 | -85.0122
6.7454 | -85.0132
6.7457 | -85.0147
6.746 | -85.0164
6.7463 | -85.018
6.7467 | -85.0193
6.747 | -85.0207
6.7473 | -85.0221
6.7477 | -85.0241
6.748 | -85.026
6.7483 | -85.0275
6.7486 | -85.0287
6.749 | -85.0304
6.7493 | -85.0325
6.7496 | -85.034
6.75 | -85.0356
6.7503 | -85.0366
6.7506 | -85.0376
6.7509 | -85.0393
6.7513 | -85.0409
6.7516 | -85.042
6.7519 | -85.043
6.7523 | -85.0445
6.7526 | -85.0459
6.7529 | -85.0475
6.7533 | -85.0492
6.7536 | -85.0504
6.7539 | -85.0525
6.7542 | -85.0566
6.7546 | -85.0595
6.7549 | -85.0608
6.7552 | -85.0622
6.7556 | -85.0635
6.7559 | -85.0647
6.7562 | -85.0658
6.7565 | -85.0671
6.7569 | -85.0684
6.7572 | -85.0698
6.7575 | -85.0714
6.7579 | -85.0727
6.7582 | -85.0741
6.7585 | -85.0757
6.7588 | -85.0773
6.7592 | -85.079
6.7595 | -85.0808
6.7598 | -85.0819
6.7602 | -85.0832
6.7605 | -85.0848
6.7608 | -85.0862
6.7612 | -85.0874
6.7615 | -85.0886
6.7618 | -85.0902
6.7621 | -85.0915
6.7625 | -85.0928
6.7628 | -85.0939
6.7631 | -85.0953
6.7635 | -85.0973
6.7638 | -85.0995
6.7641 | -85.1014
6.7644 | -85.1028
6.7648 | -85.104
6.7651 | -85.1052
6.7654 | -85.1069
6.7658 | -85.109
6.7661 | -85.1115
6.7664 | -85.1131
6.7667 | -85.1142
6.7671 | -85.116
6.7674 | -85.1169
6.7677 | -85.1179
6.7681 | -85.1196
6.7684 | -85.1213
6.7687 | -85.122
6.7691 | -85.1236
6.7694 | -85.1261
6.7697 | -85.1282
6.77 | -85.13
6.7704 | -85.1317
6.7707 | -85.1323
6.771 | -85.1343
6.7714 | -85.1363
6.7717 | -85.1381
6.772 | -85.1408
6.7723 | -85.1432
6.7727 | -85.1447
6.773 | -85.1462
6.7733 | -85.1498
6.7737 | -85.1513
6.774 | -85.1529
6.7743 | -85.1546
6.7746 | -85.1559
6.775 | -85.1577
6.7753 | -85.1589
6.7756 | -85.1609
6.776 | -85.1617
6.7763 | -85.1637
6.7766 | -85.1658
6.777 | -85.1674
6.7773 | -85.169
6.7776 | -85.1704
6.7779 | -85.1719
6.7783 | -85.1733
6.7786 | -85.1745
6.7789 | -85.1759
6.7793 | -85.1768
6.7796 | -85.1774
6.7799 | -85.1785
6.7802 | -85.18
6.7806 | -85.1816
6.7809 | -85.183
6.7812 | -85.1855
6.7816 | -85.1873
6.7819 | -85.1884
6.7822 | -85.1896
6.7825 | -85.1909
6.7829 | -85.1922
6.7832 | -85.1936
6.7835 | -85.1955
6.7839 | -85.1979
6.7842 | -85.1996
6.7845 | -85.2009
6.7848 | -85.2023
6.7852 | -85.2039
6.7855 | -85.2054
6.7858 | -85.2067
6.7862 | -85.2083
6.7865 | -85.2098
6.7868 | -85.2107
6.7872 | -85.2125
6.7875 | -85.2152
6.7878 | -85.2167
6.7881 | -85.2181
6.7885 | -85.2195
6.7888 | -85.2205
6.7891 | -85.2215
6.7895 | -85.2234
6.7898 | -85.2252
6.7901 | -85.2263
6.7904 | -85.2274
6.7908 | -85.2293
6.7911 | -85.2306
6.7914 | -85.2319
6.7918 | -85.2329
6.7921 | -85.2339
6.7924 | -85.2344
6.7927 | -85.2358
6.7931 | -85.2373
6.7934 | -85.2387
6.7937 | -85.2402
6.7941 | -85.2415
6.7944 | -85.2427
6.7947 | -85.2442
6.7951 | -85.2452
6.7954 | -85.2459
6.7957 | -85.2474
6.796 | -85.2492
6.7964 | -85.2505
6.7967 | -85.2517
6.797 | -85.2533
6.7974 | -85.2544
6.7977 | -85.2559
6.798 | -85.2579
6.7983 | -85.2589
6.7987 | -85.2601
6.799 | -85.2627
6.7993 | -85.2642
6.7997 | -85.2661
6.8 | -85.2672
6.8003 | -85.2687
6.8006 | -85.2702
6.801 | -85.2716
6.8013 | -85.2732
6.8016 | -85.2751
6.802 | -85.2777
6.8023 | -85.2797
6.8026 | -85.2808
6.803 | -85.2833
6.8033 | -85.2849
6.8036 | -85.2859
6.8039 | -85.2869
6.8043 | -85.2886
6.8046 | -85.2902
6.8049 | -85.2921
6.8053 | -85.2947
6.8056 | -85.2968
6.8059 | -85.2985
6.8062 | -85.3005
6.8066 | -85.302
6.8069 | -85.3035
6.8072 | -85.3049
6.8076 | -85.3063
6.8079 | -85.3077
6.8082 | -85.3091
6.8085 | -85.3105
6.8089 | -85.3117
6.8092 | -85.313
6.8095 | -85.3144
6.8099 | -85.3158
6.8102 | -85.3167
6.8105 | -85.318
6.8109 | -85.3193
6.8112 | -85.3206
6.8115 | -85.3222
6.8118 | -85.3235
6.8122 | -85.3258
6.8125 | -85.3284
6.8128 | -85.33
6.8132 | -85.3317
6.8135 | -85.3332
6.8138 | -85.3346
6.8141 | -85.3347
6.8145 | -85.3364
6.8148 | -85.339
6.8151 | -85.341
6.8155 | -85.3422
6.8158 | -85.3432
6.8161 | -85.344
6.8164 | -85.346
6.8168 | -85.3479
6.8171 | -85.349
6.8174 | -85.3504
6.8178 | -85.3531
6.8181 | -85.3545
6.8184 | -85.3559
6.8187 | -85.3567
6.8191 | -85.358
6.8194 | -85.3594
6.8197 | -85.3605
6.8201 | -85.3613
6.8204 | -85.3623
6.8207 | -85.3636
6.8211 | -85.3646
6.8214 | -85.3658
6.8217 | -85.367
6.822 | -85.3681
6.8224 | -85.3692
6.8227 | -85.3703
6.823 | -85.3715
6.8234 | -85.3726
6.8237 | -85.3736
6.824 | -85.3748
6.8243 | -85.3761
6.8247 | -85.3778
6.825 | -85.3798
6.8253 | -85.3818
6.8257 | -85.3832
6.826 | -85.3845
6.8263 | -85.386
6.8266 | -85.3875
6.827 | -85.3888
6.8273 | -85.39
6.8276 | -85.3916
6.828 | -85.3929
6.8283 | -85.3931
6.8286 | -85.3936
6.829 | -85.3949
6.8293 | -85.396
6.8296 | -85.3973
6.8299 | -85.3988
6.8303 | -85.3999
6.8306 | -85.4007
6.8309 | -85.4017
6.8313 | -85.4031
6.8316 | -85.4043
6.8319 | -85.4047
6.8322 | -85.4061
6.8326 | -85.4075
6.8329 | -85.4092
6.8332 | -85.412
6.8336 | -85.4158
6.8339 | -85.4196
6.8342 | -85.4228
6.8345 | -85.4249
6.8349 | -85.4264
6.8352 | -85.4276
6.8355 | -85.4282
6.8359 | -85.4293
6.8362 | -85.4305
6.8365 | -85.4316
6.8369 | -85.4319
6.8372 | -85.4332
6.8375 | -85.4339
6.8378 | -85.4351
6.8382 | -85.4365
6.8385 | -85.4381
6.8388 | -85.4398
6.8392 | -85.4411
6.8395 | -85.4426
6.8398 | -85.4433
6.8401 | -85.4444
6.8405 | -85.4457
6.8408 | -85.4469
6.8411 | -85.4485
6.8415 | -85.4506
6.8418 | -85.4524
6.8421 | -85.4536
6.8424 | -85.4548
6.8428 | -85.4552
6.8431 | -85.4566
6.8434 | -85.4571
6.8438 | -85.4586
6.8441 | -85.4606
6.8444 | -85.4618
6.8448 | -85.4629
6.8451 | -85.464
6.8454 | -85.4658
6.8457 | -85.4679
6.8461 | -85.4699
6.8464 | -85.4713
6.8467 | -85.4728
6.8471 | -85.474
6.8474 | -85.4758
6.8477 | -85.4776
6.848 | -85.4786
6.8484 | -85.4795
6.8487 | -85.4799
6.849 | -85.4805
6.8494 | -85.4819
6.8497 | -85.4836
6.85 | -85.4851
6.8503 | -85.4865
6.8507 | -85.4876
6.851 | -85.4884
6.8513 | -85.4889
6.8517 | -85.4902
6.852 | -85.4923
6.8523 | -85.4936
6.8527 | -85.4947
6.853 | -85.4967
6.8533 | -85.5002
6.8536 | -85.5016
6.854 | -85.5026
6.8543 | -85.5035
6.8546 | -85.5041
6.855 | -85.5049
6.8553 | -85.5059
6.8556 | -85.5066
6.8559 | -85.507
6.8563 | -85.5077
6.8566 | -85.5098
6.8569 | -85.5114
6.8573 | -85.5145
6.8576 | -85.5169
6.8579 | -85.518
6.8582 | -85.5187
6.8586 | -85.5197
6.8589 | -85.5216
6.8592 | -85.5228
6.8596 | -85.5234
6.8599 | -85.5246
6.8602 | -85.5252
6.8606 | -85.5257
6.8609 | -85.5278
6.8612 | -85.5299
6.8615 | -85.5309
6.8619 | -85.532
6.8622 | -85.5338
6.8625 | -85.5348
6.8629 | -85.536
6.8632 | -85.5373
6.8635 | -85.5393
6.8638 | -85.541
6.8642 | -85.542
6.8645 | -85.5423
6.8648 | -85.5433
6.8652 | -85.5446
6.8655 | -85.5456
6.8658 | -85.547
6.8661 | -85.5483
6.8665 | -85.5495
6.8668 | -85.5507
6.8671 | -85.5521
6.8675 | -85.5534
6.8678 | -85.5544
6.8681 | -85.5556
6.8685 | -85.5571
6.8688 | -85.5584
6.8691 | -85.5594
6.8694 | -85.5613
6.8698 | -85.5635
6.8701 | -85.5647
6.8704 | -85.5666
6.8708 | -85.5694
6.8711 | -85.5717
6.8714 | -85.5732
6.8717 | -85.575
6.8721 | -85.577
6.8724 | -85.5785
6.8727 | -85.5799
6.8731 | -85.5807
6.8734 | -85.5817
6.8737 | -85.5835
6.874 | -85.5851
6.8744 | -85.5865
6.8747 | -85.5888
6.875 | -85.5906
6.8754 | -85.5923
6.8757 | -85.5934
6.876 | -85.5944
6.8763 | -85.5954
6.8767 | -85.5965
6.877 | -85.5974
6.8773 | -85.5984
6.8777 | -85.5991
6.878 | -85.6007
6.8783 | -85.6025
6.8787 | -85.6042
6.879 | -85.6056
6.8793 | -85.6076
6.8796 | -85.6102
6.88 | -85.6119
6.8803 | -85.6128
6.8806 | -85.6151
6.881 | -85.6156
6.8813 | -85.6142
6.8816 | -85.614
6.8819 | -85.6151
6.8823 | -85.6172
6.8826 | -85.6187
6.8829 | -85.6198
6.8833 | -85.6222
6.8836 | -85.6238
6.8839 | -85.6269
6.8842 | -85.6283
6.8846 | -85.6293
6.8849 | -85.6305
6.8852 | -85.6316
6.8856 | -85.6326
6.8859 | -85.6336
6.8862 | -85.6347
6.8866 | -85.6354
6.8869 | -85.6362
6.8872 | -85.6381
6.8875 | -85.6402
6.8879 | -85.642
6.8882 | -85.6435
6.8885 | -85.645
6.8889 | -85.6466
6.8892 | -85.6477
6.8895 | -85.6494
6.8898 | -85.65
6.8902 | -85.6506
6.8905 | -85.6517
6.8908 | -85.6529
6.8912 | -85.6539
6.8915 | -85.6547
6.8918 | -85.656
6.8921 | -85.6574
6.8925 | -85.6584
6.8928 | -85.6592
6.8931 | -85.6601
6.8935 | -85.6614
6.8938 | -85.6624
6.8941 | -85.6638
6.8945 | -85.6651
6.8948 | -85.6666
6.8951 | -85.6678
6.8954 | -85.6693
6.8958 | -85.6708
6.8961 | -85.6719
6.8964 | -85.673
6.8968 | -85.6734
6.8971 | -85.674
6.8974 | -85.675
6.8977 | -85.6755
6.8981 | -85.6767
6.8984 | -85.6774
6.8987 | -85.6776
6.8991 | -85.6788
6.8994 | -85.6792
6.8997 | -85.6799
6.9 | -85.6817
6.9004 | -85.6827
6.9007 | -85.6842
6.901 | -85.6854
6.9014 | -85.686
6.9017 | -85.6868
6.902 | -85.6888
6.9024 | -85.6904
6.9027 | -85.6914
6.903 | -85.6926
6.9033 | -85.6942
6.9037 | -85.6957
6.904 | -85.6975
6.9043 | -85.6985
6.9047 | -85.6997
6.905 | -85.7003
6.9053 | -85.7001
6.9056 | -85.7008
6.906 | -85.7018
6.9063 | -85.7024
6.9066 | -85.7031
6.907 | -85.704
6.9073 | -85.7046
6.9076 | -85.7052
6.9079 | -85.7062
6.9083 | -85.7072
6.9086 | -85.7087
6.9089 | -85.7096
6.9093 | -85.7105
6.9096 | -85.7123
6.9099 | -85.7143
6.9103 | -85.7157
6.9106 | -85.7177
6.9109 | -85.7193
6.9112 | -85.7203
6.9116 | -85.7219
6.9119 | -85.7238
6.9122 | -85.7258
6.9126 | -85.7273
6.9129 | -85.7278
6.9132 | -85.7285
6.9135 | -85.7289
6.9139 | -85.7298
6.9142 | -85.7308
6.9145 | -85.7318
6.9149 | -85.7329
6.9152 | -85.7332
6.9155 | -85.7341
6.9158 | -85.7356
6.9162 | -85.7365
6.9165 | -85.7374
6.9168 | -85.7397
6.9172 | -85.7418
6.9175 | -85.7428
6.9178 | -85.7429
6.9181 | -85.7432
6.9185 | -85.7448
6.9188 | -85.7479
6.9191 | -85.7496
6.9195 | -85.7503
6.9198 | -85.7504
6.9201 | -85.7513
6.9205 | -85.7524
6.9208 | -85.7536
6.9211 | -85.7546
6.9214 | -85.7556
6.9218 | -85.7565
6.9221 | -85.7574
6.9224 | -85.7582
6.9228 | -85.7581
6.9231 | -85.7591
6.9234 | -85.76
6.9237 | -85.7608
6.9241 | -85.7615
6.9244 | -85.761
6.9247 | -85.7617
6.9251 | -85.7628
6.9254 | -85.7638
6.9257 | -85.7647
6.926 | -85.7659
6.9264 | -85.7672
6.9267 | -85.768
6.927 | -85.769
6.9274 | -85.77
6.9277 | -85.7709
6.928 | -85.7719
6.9284 | -85.7729
6.9287 | -85.7738
6.929 | -85.7749
6.9293 | -85.7771
6.9297 | -85.7786
6.93 | -85.7793
6.9303 | -85.78
6.9307 | -85.7808
6.931 | -85.7814
6.9313 | -85.7816
6.9316 | -85.7827
6.932 | -85.7842
6.9323 | -85.7858
6.9326 | -85.7875
6.933 | -85.7888
6.9333 | -85.7905
6.9336 | -85.7911
6.9339 | -85.7921
6.9343 | -85.7941
6.9346 | -85.7952
6.9349 | -85.7963
6.9353 | -85.7974
6.9356 | -85.7986
6.9359 | -85.8
6.9363 | -85.8012
6.9366 | -85.8023
6.9369 | -85.8033
6.9372 | -85.8049
6.9376 | -85.8059
6.9379 | -85.8068
6.9382 | -85.8078
6.9386 | -85.8081
6.9389 | -85.8091
6.9392 | -85.8096
6.9395 | -85.8104
6.9399 | -85.8116
6.9402 | -85.8127
6.9405 | -85.8145
6.9409 | -85.8159
6.9412 | -85.8163
6.9415 | -85.8172
6.9418 | -85.8185
6.9422 | -85.8193
6.9425 | -85.8201
6.9428 | -85.8209
6.9432 | -85.822
6.9435 | -85.8231
6.9438 | -85.8241
6.9442 | -85.825
6.9445 | -85.8259
6.9448 | -85.827
6.9451 | -85.8279
6.9455 | -85.8286
6.9458 | -85.83
6.9461 | -85.8325
6.9465 | -85.8342
6.9468 | -85.8353
6.9471 | -85.8372
6.9474 | -85.8387
6.9478 | -85.8395
6.9481 | -85.8404
6.9484 | -85.8411
6.9488 | -85.8427
6.9491 | -85.8442
6.9494 | -85.8451
6.9497 | -85.8456
6.9501 | -85.8466
6.9504 | -85.8477
6.9507 | -85.8484
6.9511 | -85.8477
6.9514 | -85.8487
6.9517 | -85.8499
6.9521 | -85.8515
6.9524 | -85.8531
6.9527 | -85.8543
6.953 | -85.8559
6.9534 | -85.8571
6.9537 | -85.8585
6.954 | -85.8596
6.9544 | -85.8606
6.9547 | -85.8615
6.955 | -85.8626
6.9553 | -85.8637
6.9557 | -85.8644
6.956 | -85.8656
6.9563 | -85.8668
6.9567 | -85.8678
6.957 | -85.8682
6.9573 | -85.868
6.9576 | -85.8691
6.958 | -85.871
6.9583 | -85.8733
6.9586 | -85.8751
6.959 | -85.8758
6.9593 | -85.8765
6.9596 | -85.8776
6.96 | -85.8779
6.9603 | -85.8795
6.9606 | -85.8805
6.9609 | -85.8813
6.9613 | -85.8818
6.9616 | -85.8824
6.9619 | -85.8841
6.9623 | -85.8856
6.9626 | -85.8861
6.9629 | -85.8871
6.9632 | -85.8885
6.9636 | -85.89
6.9639 | -85.8907
6.9642 | -85.8916
6.9646 | -85.8924
6.9649 | -85.893
6.9652 | -85.894
6.9655 | -85.8946
6.9659 | -85.895
6.9662 | -85.8955
6.9665 | -85.8962
6.9669 | -85.8973
6.9672 | -85.8983
6.9675 | -85.8995
6.9679 | -85.9007
6.9682 | -85.9019
6.9685 | -85.9034
6.9688 | -85.9041
6.9692 | -85.9044
6.9695 | -85.905
6.9698 | -85.9063
6.9702 | -85.9063
6.9705 | -85.9074
6.9708 | -85.9087
6.9711 | -85.9097
6.9715 | -85.9112
6.9718 | -85.9122
6.9721 | -85.914
6.9725 | -85.9166
6.9728 | -85.9175
6.9731 | -85.9208
6.9734 | -85.9219
6.9738 | -85.9226
6.9741 | -85.9236
6.9744 | -85.9246
6.9748 | -85.9254
6.9751 | -85.9264
6.9754 | -85.9274
6.9757 | -85.9284
6.9761 | -85.9295
6.9764 | -85.9305
6.9767 | -85.9317
6.9771 | -85.9321
6.9774 | -85.9326
6.9777 | -85.9335
6.9781 | -85.9345
6.9784 | -85.9355
6.9787 | -85.9361
6.979 | -85.9368
6.9794 | -85.9382
6.9797 | -85.9392
6.98 | -85.9413
6.9804 | -85.9424
6.9807 | -85.9433
6.981 | -85.9449
6.9813 | -85.9467
6.9817 | -85.948
6.982 | -85.949
6.9823 | -85.9497
6.9827 | -85.951
6.983 | -85.9521
6.9833 | -85.9524
6.9836 | -85.9529
6.984 | -85.954
6.9843 | -85.9548
6.9846 | -85.9557
6.985 | -85.9564
6.9853 | -85.9567
6.9856 | -85.9572
6.986 | -85.9592
6.9863 | -85.9615
6.9866 | -85.9624
6.9869 | -85.9635
6.9873 | -85.9644
6.9876 | -85.9651
6.9879 | -85.9662
6.9883 | -85.9681
6.9886 | -85.9697
6.9889 | -85.9709
6.9892 | -85.9723
6.9896 | -85.9731
6.9899 | -85.9743
6.9902 | -85.9754
6.9906 | -85.9762
6.9909 | -85.9771
6.9912 | -85.9776
6.9915 | -85.979
6.9919 | -85.9803
6.9922 | -85.981
6.9925 | -85.9817
6.9929 | -85.9827
6.9932 | -85.984
6.9935 | -85.9847
6.9939 | -85.9851
6.9942 | -85.9861
6.9945 | -85.987
6.9948 | -85.9886
6.9952 | -85.9898
6.9955 | -85.9906
6.9958 | -85.9921
6.9962 | -85.9931
6.9965 | -85.9938
6.9968 | -85.9945
6.9971 | -85.9953
6.9975 | -85.9961
6.9978 | -85.9971
6.9981 | -85.9985
6.9985 | -85.9998
6.9988 | -86.0005
6.9991 | -86.0006
6.9994 | -86.0013
6.9998 | -86.0023
7.0001 | -86.0029
7.0004 | -86.004
7.0008 | -86.0055
7.0011 | -86.0063
7.0014 | -86.0072
7.0018 | -86.0085
7.0021 | -86.0098
7.0024 | -86.0105
7.0027 | -86.0117
7.0031 | -86.0121
7.0034 | -86.0123
7.0037 | -86.0134
7.0041 | -86.0144
7.0044 | -86.0155
7.0047 | -86.0164
7.005 | -86.0174
7.0054 | -86.0188
7.0057 | -86.0201
7.006 | -86.0208
7.0064 | -86.0216
7.0067 | -86.0224
7.007 | -86.0232
7.0073 | -86.0235
7.0077 | -86.024
7.008 | -86.0254
7.0083 | -86.0245
7.0087 | -86.025
7.009 | -86.0257
7.0093 | -86.0257
7.0096 | -86.0245
7.01 | -86.0249
7.0103 | -86.0258
7.0106 | -86.0263
7.011 | -86.0271
7.0113 | -86.0284
7.0116 | -86.0297
7.012 | -86.0307
7.0123 | -86.0313
7.0126 | -86.0321
7.0129 | -86.033
7.0133 | -86.033
7.0136 | -86.0327
7.0139 | -86.0331
7.0143 | -86.034
7.0146 | -86.035
7.0149 | -86.0354
7.0152 | -86.0349
7.0156 | -86.0347
7.0159 | -86.0355
7.0162 | -86.0364
7.0166 | -86.0375
7.0169 | -86.0397
7.0172 | -86.041
7.0175 | -86.0418
7.0179 | -86.043
7.0182 | -86.0443
7.0185 | -86.0451
7.0189 | -86.0457
7.0192 | -86.0469
7.0195 | -86.0491
7.0199 | -86.0514
7.0202 | -86.0532
7.0205 | -86.0547
7.0208 | -86.0554
7.0212 | -86.0572
7.0215 | -86.06
7.0218 | -86.062
7.0222 | -86.0631
7.0225 | -86.0642
7.0228 | -86.0651
7.0231 | -86.0658
7.0235 | -86.0667
7.0238 | -86.0674
7.0241 | -86.0681
7.0245 | -86.0688
7.0248 | -86.0696
7.0251 | -86.0708
7.0254 | -86.072
7.0258 | -86.0735
7.0261 | -86.074
7.0264 | -86.0738
7.0268 | -86.0745
7.0271 | -86.0756
7.0274 | -86.0759
7.0278 | -86.0763
7.0281 | -86.0769
7.0284 | -86.0775
7.0287 | -86.0788
7.0291 | -86.0803
7.0294 | -86.0821
7.0297 | -86.0828
7.0301 | -86.0841
7.0304 | -86.0843
7.0307 | -86.0864
7.031 | -86.0872
7.0314 | -86.088
7.0317 | -86.0891
7.032 | -86.0902
7.0324 | -86.0917
7.0327 | -86.093
7.033 | -86.0945
7.0333 | -86.0948
7.0337 | -86.0952
7.034 | -86.0958
7.0343 | -86.0949
7.0347 | -86.0943
7.035 | -86.0947
7.0353 | -86.0956
7.0357 | -86.0964
7.036 | -86.0972
7.0363 | -86.0983
7.0366 | -86.0993
7.037 | -86.1004
7.0373 | -86.1017
7.0376 | -86.1026
7.038 | -86.1036
7.0383 | -86.1051
7.0386 | -86.1055
7.0389 | -86.1064
7.0393 | -86.1077
7.0396 | -86.1086
7.0399 | -86.1094
7.0403 | -86.1101
7.0406 | -86.1105
7.0409 | -86.1105
7.0412 | -86.1111
7.0416 | -86.1115
7.0419 | -86.1118
7.0422 | -86.1133
7.0426 | -86.1154
7.0429 | -86.1163
7.0432 | -86.117
7.0436 | -86.1175
7.0439 | -86.1182
7.0442 | -86.118
7.0445 | -86.1184
7.0449 | -86.1193
7.0452 | -86.1201
7.0455 | -86.1206
7.0459 | -86.12
7.0462 | -86.1201
7.0465 | -86.1193
7.0468 | -86.1195
7.0472 | -86.1204
7.0475 | -86.1212
7.0478 | -86.1224
7.0482 | -86.1235
7.0485 | -86.1245
7.0488 | -86.1251
7.0491 | -86.1257
7.0495 | -86.1262
7.0498 | -86.1265
7.0501 | -86.1272
7.0505 | -86.1274
7.0508 | -86.1275
7.0511 | -86.1283
7.0515 | -86.1289
7.0518 | -86.1298
7.0521 | -86.1318
7.0524 | -86.1337
7.0528 | -86.1349
7.0531 | -86.1356
7.0534 | -86.1363
7.0538 | -86.1371
7.0541 | -86.1378
7.0544 | -86.1379
7.0547 | -86.1386
7.0551 | -86.1393
7.0554 | -86.1392
7.0557 | -86.1398
7.0561 | -86.1406
7.0564 | -86.1403
7.0567 | -86.139
7.057 | -86.1393
7.0574 | -86.139
7.0577 | -86.1393
7.058 | -86.1408
7.0584 | -86.142
7.0587 | -86.1424
7.059 | -86.143
7.0594 | -86.1438
7.0597 | -86.1439
7.06 | -86.1447
7.0603 | -86.1456
7.0607 | -86.1465
7.061 | -86.1471
7.0613 | -86.148
7.0617 | -86.1476
7.062 | -86.1471
7.0623 | -86.1478
7.0626 | -86.1489
7.063 | -86.1496
7.0633 | -86.1507
7.0636 | -86.1519
7.064 | -86.1529
7.0643 | -86.1539
7.0646 | -86.1549
7.0649 | -86.1558
7.0653 | -86.1561
7.0656 | -86.1571
7.0659 | -86.1577
7.0663 | -86.1577
7.0666 | -86.1583
7.0669 | -86.1591
7.0672 | -86.1597
7.0676 | -86.1596
7.0679 | -86.1592
7.0682 | -86.1601
7.0686 | -86.1609
7.0689 | -86.1612
7.0692 | -86.1616
7.0696 | -86.1619
7.0699 | -86.1622
7.0702 | -86.1632
7.0705 | -86.1639
7.0709 | -86.1638
7.0712 | -86.1642
7.0715 | -86.1646
7.0719 | -86.1652
7.0722 | -86.1658
7.0725 | -86.1662
7.0728 | -86.1668
7.0732 | -86.1674
7.0735 | -86.1685
7.0738 | -86.1693
7.0742 | -86.1699
7.0745 | -86.1706
7.0748 | -86.1707
7.0751 | -86.171
7.0755 | -86.1714
7.0758 | -86.1716
7.0761 | -86.1721
7.0765 | -86.1725
7.0768 | -86.173
7.0771 | -86.1735
7.0775 | -86.1741
7.0778 | -86.1738
7.0781 | -86.1732
7.0784 | -86.1736
7.0788 | -86.1749
7.0791 | -86.1753
7.0794 | -86.176
7.0798 | -86.1764
7.0801 | -86.1768
7.0804 | -86.177
7.0807 | -86.1771
7.0811 | -86.1768
7.0814 | -86.1773
7.0817 | -86.1782
7.0821 | -86.1785
7.0824 | -86.1786
7.0827 | -86.1798
7.083 | -86.1809
7.0834 | -86.1806
7.0837 | -86.1803
7.084 | -86.1808
7.0844 | -86.1813
7.0847 | -86.1819
7.085 | -86.183
7.0854 | -86.1838
7.0857 | -86.1841
7.086 | -86.1847
7.0863 | -86.1853
7.0867 | -86.1859
7.087 | -86.1866
7.0873 | -86.1873
7.0877 | -86.188
7.088 | -86.1884
7.0883 | -86.1894
7.0886 | -86.1914
7.089 | -86.1924
7.0893 | -86.1932
7.0896 | -86.1933
7.09 | -86.1955
7.0903 | -86.1956
7.0906 | -86.196
7.0909 | -86.1961
7.0913 | -86.1962
7.0916 | -86.1968
7.0919 | -86.1973
7.0923 | -86.1975
7.0926 | -86.1978
7.0929 | -86.1985
7.0933 | -86.1991
7.0936 | -86.1994
7.0939 | -86.1994
7.0942 | -86.1996
7.0946 | -86.2003
7.0949 | -86.2014
7.0952 | -86.2021
7.0956 | -86.2027
7.0959 | -86.2032
7.0962 | -86.2044
7.0965 | -86.2048
7.0969 | -86.205
7.0972 | -86.2048
7.0975 | -86.2053
7.0979 | -86.2059
7.0982 | -86.2063
7.0985 | -86.2064
7.0988 | -86.2065
7.0992 | -86.207
7.0995 | -86.2064
7.0998 | -86.2067
7.1002 | -86.2072
7.1005 | -86.2075
7.1008 | -86.208
7.1012 | -86.2087
7.1015 | -86.2093
7.1018 | -86.2098
7.1021 | -86.2104
7.1025 | -86.2113
7.1028 | -86.2122
7.1031 | -86.2123
7.1035 | -86.2139
7.1038 | -86.2159
7.1041 | -86.2177
7.1044 | -86.2178
7.1048 | -86.2179
7.1051 | -86.2183
7.1054 | -86.2186
7.1058 | -86.2186
7.1061 | -86.2189
7.1064 | -86.2194
7.1067 | -86.2192
7.1071 | -86.2196
7.1074 | -86.2199
7.1077 | -86.2205
7.1081 | -86.2217
7.1084 | -86.2227
7.1087 | -86.2231
7.109 | -86.2236
7.1094 | -86.2233
7.1097 | -86.2234
7.11 | -86.2241
7.1104 | -86.2241
7.1107 | -86.2242
7.111 | -86.2245
7.1114 | -86.225
7.1117 | -86.2254
7.112 | -86.2257
7.1123 | -86.2265
7.1127 | -86.2263
7.113 | -86.2272
7.1133 | -86.2274
7.1137 | -86.2266
7.114 | -86.227
7.1143 | -86.2274
7.1146 | -86.2275
7.115 | -86.2278
7.1153 | -86.2277
7.1156 | -86.2268
7.116 | -86.2273
7.1163 | -86.2284
7.1166 | -86.2294
7.1169 | -86.23
7.1173 | -86.2303
7.1176 | -86.2306
7.1179 | -86.2304
7.1183 | -86.2308
7.1186 | -86.2333
7.1189 | -86.234
7.1193 | -86.2343
7.1196 | -86.235
7.1199 | -86.2359
7.1202 | -86.2364
7.1206 | -86.2368
7.1209 | -86.2361
7.1212 | -86.2352
7.1216 | -86.2352
7.1219 | -86.2354
7.1222 | -86.2351
7.1225 | -86.2352
7.1229 | -86.2354
7.1232 | -86.2359
7.1235 | -86.2361
7.1239 | -86.2364
7.1242 | -86.2368
7.1245 | -86.2382
7.1248 | -86.2389
7.1252 | -86.239
7.1255 | -86.2394
7.1258 | -86.2402
7.1262 | -86.2399
7.1265 | -86.2401
7.1268 | -86.2408
7.1272 | -86.242
7.1275 | -86.2435
7.1278 | -86.2439
7.1281 | -86.2444
7.1285 | -86.2465
7.1288 | -86.2475
7.1291 | -86.2476
7.1295 | -86.2477
7.1298 | -86.2481
7.1301 | -86.2483
7.1304 | -86.249
7.1308 | -86.2505
7.1311 | -86.2518
7.1314 | -86.253
7.1318 | -86.2537
7.1321 | -86.2548
7.1324 | -86.2552
7.1327 | -86.2557
7.1331 | -86.2568
7.1334 | -86.2571
7.1337 | -86.2574
7.1341 | -86.2577
7.1344 | -86.2577
7.1347 | -86.2575
7.1351 | -86.2577
7.1354 | -86.2579
7.1357 | -86.2585
7.136 | -86.2586
7.1364 | -86.2582
7.1367 | -86.258
7.137 | -86.2584
7.1374 | -86.2595
7.1377 | -86.2603
7.138 | -86.2605
7.1383 | -86.261
7.1387 | -86.2613
7.139 | -86.2611
7.1393 | -86.261
7.1397 | -86.2617
7.14 | -86.2619
7.1403 | -86.2624
7.1406 | -86.2629
7.141 | -86.2633
7.1413 | -86.2639
7.1416 | -86.2647
7.142 | -86.2663
7.1423 | -86.2659
7.1426 | -86.2651
7.143 | -86.265
7.1433 | -86.2651
7.1436 | -86.2664
7.1439 | -86.2668
7.1443 | -86.2672
7.1446 | -86.2676
7.1449 | -86.2682
7.1453 | -86.2685
7.1456 | -86.2694
7.1459 | -86.2702
7.1462 | -86.2707
7.1466 | -86.271
7.1469 | -86.2712
7.1472 | -86.2718
7.1476 | -86.2721
7.1479 | -86.2724
7.1482 | -86.2726
7.1485 | -86.2727
7.1489 | -86.2728
7.1492 | -86.2729
7.1495 | -86.2736
7.1499 | -86.2741
7.1502 | -86.274
7.1505 | -86.2744
7.1509 | -86.2742
7.1512 | -86.2744
7.1515 | -86.2746
7.1518 | -86.2748
7.1522 | -86.275
7.1525 | -86.2753
7.1528 | -86.2757
7.1532 | -86.2767
7.1535 | -86.2768
7.1538 | -86.2769
7.1541 | -86.2776
7.1545 | -86.2782
7.1548 | -86.2786
7.1551 | -86.2788
7.1555 | -86.2791
7.1558 | -86.2791
7.1561 | -86.2796
7.1564 | -86.2801
7.1568 | -86.2814
7.1571 | -86.2819
7.1574 | -86.2825
7.1578 | -86.2833
7.1581 | -86.2844
7.1584 | -86.2847
7.1588 | -86.2845
7.1591 | -86.2847
7.1594 | -86.2853
7.1597 | -86.2855
7.1601 | -86.2858
7.1604 | -86.2857
7.1607 | -86.2858
7.1611 | -86.2856
7.1614 | -86.2852
7.1617 | -86.2854
7.162 | -86.2857
7.1624 | -86.2858
7.1627 | -86.2863
7.163 | -86.2869
7.1634 | -86.2875
7.1637 | -86.2877
7.164 | -86.288
7.1643 | -86.2884
7.1647 | -86.2889
7.165 | -86.2892
7.1653 | -86.2896
7.1657 | -86.2895
7.166 | -86.2897
7.1663 | -86.2898
7.1666 | -86.2897
7.167 | -86.2897
7.1673 | -86.2903
7.1676 | -86.2917
7.168 | -86.2923
7.1683 | -86.2925
7.1686 | -86.2929
7.169 | -86.2932
7.1693 | -86.2935
7.1696 | -86.2942
7.1699 | -86.2936
7.1703 | -86.2935
7.1706 | -86.2939
7.1709 | -86.2947
7.1713 | -86.2961
7.1716 | -86.2966
7.1719 | -86.2963
7.1722 | -86.2965
7.1726 | -86.2969
7.1729 | -86.2967
7.1732 | -86.2969
7.1736 | -86.2969
7.1739 | -86.2974
7.1742 | -86.2979
7.1745 | -86.2983
7.1749 | -86.2987
7.1752 | -86.299
7.1755 | -86.2993
7.1759 | -86.299
7.1762 | -86.2992
7.1765 | -86.3002
7.1769 | -86.3008
7.1772 | -86.3005
7.1775 | -86.3003
7.1778 | -86.3006
7.1782 | -86.3013
7.1785 | -86.3023
7.1788 | -86.3027
7.1792 | -86.3032
7.1795 | -86.3036
7.1798 | -86.304
7.1801 | -86.3049
7.1805 | -86.305
7.1808 | -86.3053
7.1811 | -86.3052
7.1815 | -86.3054
7.1818 | -86.306
7.1821 | -86.3065
7.1824 | -86.3065
7.1828 | -86.3065
7.1831 | -86.307
7.1834 | -86.3084
7.1838 | -86.309
7.1841 | -86.3089
7.1844 | -86.3092
7.1848 | -86.31
7.1851 | -86.3096
7.1854 | -86.3089
7.1857 | -86.309
7.1861 | -86.3096
7.1864 | -86.3097
7.1867 | -86.3096
7.1871 | -86.3098
7.1874 | -86.3106
7.1877 | -86.3109
7.188 | -86.3109
7.1884 | -86.3102
7.1887 | -86.3098
7.189 | -86.3108
7.1894 | -86.3104
7.1897 | -86.3105
7.19 | -86.3108
7.1903 | -86.3113
7.1907 | -86.3112
7.191 | -86.3112
7.1913 | -86.3115
7.1917 | -86.3118
7.192 | -86.3119
7.1923 | -86.3117
7.1927 | -86.3117
7.193 | -86.3118
7.1933 | -86.313
7.1936 | -86.315
7.194 | -86.3165
7.1943 | -86.3177
7.1946 | -86.3181
7.195 | -86.3193
7.1953 | -86.3201
7.1956 | -86.3207
7.1959 | -86.3207
7.1963 | -86.3208
7.1966 | -86.3207
7.1969 | -86.3208
7.1973 | -86.3212
7.1976 | -86.3214
7.1979 | -86.3217
7.1982 | -86.322
7.1986 | -86.3218
7.1989 | -86.3217
7.1992 | -86.3218
7.1996 | -86.3222
7.1999 | -86.3223
7.2002 | -86.3217
7.2005 | -86.3211
7.2009 | -86.3215
7.2012 | -86.322
7.2015 | -86.3221
7.2019 | -86.3231
7.2022 | -86.3248
7.2025 | -86.3256
7.2029 | -86.3258
7.2032 | -86.326
7.2035 | -86.326
7.2038 | -86.3263
7.2042 | -86.3268
7.2045 | -86.3271
7.2048 | -86.3275
7.2052 | -86.3275
7.2055 | -86.3273
7.2058 | -86.3277
7.2061 | -86.3278
7.2065 | -86.3282
7.2068 | -86.3282
7.2071 | -86.3269
7.2075 | -86.3254
7.2078 | -86.326
7.2081 | -86.3262
7.2085 | -86.3266
7.2088 | -86.3273
7.2091 | -86.3275
7.2094 | -86.3278
7.2098 | -86.3282
7.2101 | -86.3284
7.2104 | -86.329
7.2108 | -86.3302
7.2111 | -86.3312
7.2114 | -86.3316
7.2117 | -86.3319
7.2121 | -86.3327
7.2124 | -86.3341
7.2127 | -86.3342
7.2131 | -86.3348
7.2134 | -86.3361
7.2137 | -86.3365
7.214 | -86.3369
7.2144 | -86.3369
7.2147 | -86.3371
7.215 | -86.3367
7.2154 | -86.3369
7.2157 | -86.3371
7.216 | -86.3371
7.2163 | -86.3373
7.2167 | -86.3376
7.217 | -86.3385
7.2173 | -86.3387
7.2177 | -86.3388
7.218 | -86.339
7.2183 | -86.3397
7.2187 | -86.3404
7.219 | -86.3413
7.2193 | -86.3416
7.2196 | -86.3422
7.22 | -86.3431
7.2203 | -86.3439
7.2206 | -86.3446
7.221 | -86.3446
7.2213 | -86.3449
7.2216 | -86.3451
7.2219 | -86.3453
7.2223 | -86.3456
7.2226 | -86.3462
7.2229 | -86.3469
7.2233 | -86.3473
7.2236 | -86.3478
7.2239 | -86.3486
7.2242 | -86.3491
7.2246 | -86.3496
7.2249 | -86.3496
7.2252 | -86.3503
7.2256 | -86.3509
7.2259 | -86.3511
7.2262 | -86.3517
7.2266 | -86.3525
7.2269 | -86.3526
7.2272 | -86.3542
7.2275 | -86.356
7.2279 | -86.357
7.2282 | -86.3579
7.2285 | -86.3577
7.2289 | -86.3579
7.2292 | -86.3577
7.2295 | -86.3578
7.2298 | -86.3582
7.2302 | -86.3585
7.2305 | -86.3593
7.2308 | -86.36
7.2312 | -86.3598
7.2315 | -86.3595
7.2318 | -86.3598
7.2321 | -86.3606
7.2325 | -86.3612
7.2328 | -86.3614
7.2331 | -86.3618
7.2335 | -86.3622
7.2338 | -86.3627
7.2341 | -86.3629
7.2345 | -86.3632
7.2348 | -86.3636
7.2351 | -86.3638
7.2354 | -86.3645
7.2358 | -86.3651
7.2361 | -86.3659
7.2364 | -86.366
7.2368 | -86.366
7.2371 | -86.3661
7.2374 | -86.3667
7.2377 | -86.3672
7.2381 | -86.3673
7.2384 | -86.3673
7.2387 | -86.3678
7.2391 | -86.3689
7.2394 | -86.3696
7.2397 | -86.3698
7.24 | -86.3701
7.2404 | -86.37
7.2407 | -86.3699
7.241 | -86.3703
7.2414 | -86.3709
7.2417 | -86.3715
7.242 | -86.3724
7.2424 | -86.3726
7.2427 | -86.3727
7.243 | -86.3728
7.2433 | -86.3725
7.2437 | -86.3726
7.244 | -86.373
7.2443 | -86.3734
7.2447 | -86.3737
7.245 | -86.3741
7.2453 | -86.3746
7.2456 | -86.3757
7.246 | -86.3763
7.2463 | -86.3769
7.2466 | -86.3775
7.247 | -86.3781
7.2473 | -86.3783
7.2476 | -86.3794
7.2479 | -86.3806
7.2483 | -86.381
7.2486 | -86.3815
7.2489 | -86.3818
7.2493 | -86.3821
7.2496 | -86.3824
7.2499 | -86.3828
7.2503 | -86.3831
7.2506 | -86.3836
7.2509 | -86.3842
7.2512 | -86.3848
7.2516 | -86.3857
7.2519 | -86.3858
7.2522 | -86.3858
7.2526 | -86.3863
7.2529 | -86.3868
7.2532 | -86.3878
7.2535 | -86.388
7.2539 | -86.389
7.2542 | -86.3904
7.2545 | -86.3905
7.2549 | -86.3909
7.2552 | -86.3921
7.2555 | -86.3935
7.2558 | -86.3943
7.2562 | -86.3949
7.2565 | -86.3954
7.2568 | -86.3958
7.2572 | -86.3965
7.2575 | -86.398
7.2578 | -86.399
7.2581 | -86.3995
7.2585 | -86.3999
7.2588 | -86.4004
7.2591 | -86.4008
7.2595 | -86.4013
7.2598 | -86.4016
7.2601 | -86.4021
7.2605 | -86.4024
7.2608 | -86.4026
7.2611 | -86.403
7.2614 | -86.4034
7.2618 | -86.4039
7.2621 | -86.4042
7.2624 | -86.405
7.2628 | -86.4056
7.2631 | -86.4066
7.2634 | -86.4071
7.2637 | -86.4073
7.2641 | -86.408
7.2644 | -86.4086
7.2647 | -86.4094
7.2651 | -86.4087
7.2654 | -86.4084
7.2657 | -86.4087
7.266 | -86.4091
7.2664 | -86.4094
7.2667 | -86.4103
7.267 | -86.4124
7.2674 | -86.4138
7.2677 | -86.4147
7.268 | -86.415
7.2684 | -86.4155
7.2687 | -86.4162
7.269 | -86.4168
7.2693 | -86.4175
7.2697 | -86.4183
7.27 | -86.4191
7.2703 | -86.4196
7.2707 | -86.4201
7.271 | -86.4207
7.2713 | -86.4212
7.2716 | -86.4217
7.272 | -86.4226
7.2723 | -86.4234
7.2726 | -86.4232
7.273 | -86.4235
7.2733 | -86.4243
7.2736 | -86.4247
7.2739 | -86.4257
7.2743 | -86.4265
7.2746 | -86.4277
7.2749 | -86.4293
7.2753 | -86.4304
7.2756 | -86.4312
7.2759 | -86.4317
7.2763 | -86.432
7.2766 | -86.4326
7.2769 | -86.4334
7.2772 | -86.4341
7.2776 | -86.4347
7.2779 | -86.4352
7.2782 | -86.4361
7.2786 | -86.4372
7.2789 | -86.4382
7.2792 | -86.4387
7.2795 | -86.4392
7.2799 | -86.4401
7.2802 | -86.4407
7.2805 | -86.4404
7.2809 | -86.4418
7.2812 | -86.4425
7.2815 | -86.4433
7.2818 | -86.4435
7.2822 | -86.4432
7.2825 | -86.4434
7.2828 | -86.4436
7.2832 | -86.4442
7.2835 | -86.4457
7.2838 | -86.447
7.2842 | -86.4483
7.2845 | -86.4487
7.2848 | -86.449
7.2851 | -86.4494
7.2855 | -86.4498
7.2858 | -86.4506
7.2861 | -86.4513
7.2865 | -86.4516
7.2868 | -86.4526
7.2871 | -86.4539
7.2874 | -86.4554
7.2878 | -86.4563
7.2881 | -86.4571
7.2884 | -86.458
7.2888 | -86.4584
7.2891 | -86.4592
7.2894 | -86.4601
7.2897 | -86.4607
7.2901 | -86.4615
7.2904 | -86.462
7.2907 | -86.463
7.2911 | -86.4639
7.2914 | -86.4647
7.2917 | -86.466
7.2921 | -86.4665
7.2924 | -86.4673
7.2927 | -86.4678
7.293 | -86.4684
7.2934 | -86.4696
7.2937 | -86.4698
7.294 | -86.4706
7.2944 | -86.4712
7.2947 | -86.4712
7.295 | -86.4712
7.2953 | -86.4703
7.2957 | -86.4693
7.296 | -86.4691
7.2963 | -86.4689
7.2967 | -86.4698
7.297 | -86.4717
7.2973 | -86.4717
7.2976 | -86.4721
7.298 | -86.4728
7.2983 | -86.4739
7.2986 | -86.4748
7.299 | -86.4753
7.2993 | -86.4752
7.2996 | -86.4741
7.2999 | -86.4735
7.3003 | -86.4741
7.3006 | -86.4752
7.3009 | -86.4754
7.3013 | -86.4752
7.3016 | -86.4754
7.3019 | -86.4762
7.3023 | -86.4768
7.3026 | -86.4781
7.3029 | -86.479
7.3032 | -86.4801
7.3036 | -86.4813
7.3039 | -86.4819
7.3042 | -86.4836
7.3046 | -86.4849
7.3049 | -86.4856
7.3052 | -86.4864
7.3055 | -86.4876
7.3059 | -86.489
7.3062 | -86.4899
7.3065 | -86.4902
7.3069 | -86.4907
7.3072 | -86.4915
7.3075 | -86.4922
7.3078 | -86.4928
7.3082 | -86.4934
7.3085 | -86.4936
7.3088 | -86.4944
7.3092 | -86.4953
7.3095 | -86.4958
7.3098 | -86.496
7.3102 | -86.4971
7.3105 | -86.4987
7.3108 | -86.4998
7.3111 | -86.5006
7.3115 | -86.5015
7.3118 | -86.5024
7.3121 | -86.5036
7.3125 | -86.5041
7.3128 | -86.5047
7.3131 | -86.5056
7.3134 | -86.5066
7.3138 | -86.5074
7.3141 | -86.5082
7.3144 | -86.5087
7.3148 | -86.5085
7.3151 | -86.5086
7.3154 | -86.5089
7.3157 | -86.5094
7.3161 | -86.5097
7.3164 | -86.5102
7.3167 | -86.511
7.3171 | -86.5117
7.3174 | -86.5125
7.3177 | -86.5133
7.3181 | -86.514
7.3184 | -86.5148
7.3187 | -86.5157
7.319 | -86.517
7.3194 | -86.5191
7.3197 | -86.5205
7.32 | -86.5207
7.3204 | -86.5217
7.3207 | -86.5225
7.321 | -86.5232
7.3213 | -86.5242
7.3217 | -86.5249
7.322 | -86.5254
7.3223 | -86.5264
7.3227 | -86.5277
7.323 | -86.5294
7.3233 | -86.5316
7.3236 | -86.5326
7.324 | -86.5323
7.3243 | -86.5322
7.3246 | -86.5329
7.325 | -86.5335
7.3253 | -86.5343
7.3256 | -86.5349
7.326 | -86.5358
7.3263 | -86.5365
7.3266 | -86.5372
7.3269 | -86.5386
7.3273 | -86.5395
7.3276 | -86.5402
7.3279 | -86.5412
7.3283 | -86.5427
7.3286 | -86.544
7.3289 | -86.5449
7.3292 | -86.546
7.3296 | -86.5469
7.3299 | -86.5477
7.3302 | -86.5485
7.3306 | -86.5493
7.3309 | -86.5501
7.3312 | -86.5514
7.3315 | -86.5525
7.3319 | -86.5538
7.3322 | -86.5548
7.3325 | -86.5559
7.3329 | -86.5567
7.3332 | -86.5575
7.3335 | -86.5587
7.3339 | -86.5597
7.3342 | -86.5605
7.3345 | -86.562
7.3348 | -86.5639
7.3352 | -86.5662
7.3355 | -86.5671
7.3358 | -86.5676
7.3362 | -86.5684
7.3365 | -86.5692
7.3368 | -86.5697
7.3371 | -86.5706
7.3375 | -86.5714
7.3378 | -86.5722
7.3381 | -86.5736
7.3385 | -86.575
7.3388 | -86.5757
7.3391 | -86.5767
7.3394 | -86.5778
7.3398 | -86.5785
7.3401 | -86.5793
7.3404 | -86.5803
7.3408 | -86.5807
7.3411 | -86.5817
7.3414 | -86.5827
7.3418 | -86.5835
7.3421 | -86.5843
7.3424 | -86.5851
7.3427 | -86.5857
7.3431 | -86.5866
7.3434 | -86.5876
7.3437 | -86.589
7.3441 | -86.5907
7.3444 | -86.5916
7.3447 | -86.5921
7.345 | -86.5926
7.3454 | -86.5934
7.3457 | -86.5943
7.346 | -86.5951
7.3464 | -86.5956
7.3467 | -86.5962
7.347 | -86.5969
7.3473 | -86.5984
7.3477 | -86.5995
7.348 | -86.6003
7.3483 | -86.6015
7.3487 | -86.6033
7.349 | -86.6047
7.3493 | -86.6055
7.3496 | -86.6061
7.35 | -86.6075
7.3503 | -86.6081
7.3506 | -86.609
7.351 | -86.6099
7.3513 | -86.6111
7.3516 | -86.6129
7.352 | -86.6145
7.3523 | -86.6156
7.3526 | -86.6172
7.3529 | -86.6189
7.3533 | -86.6198
7.3536 | -86.6212
7.3539 | -86.6229
7.3543 | -86.6239
7.3546 | -86.6247
7.3549 | -86.6258
7.3552 | -86.6273
7.3556 | -86.6285
7.3559 | -86.6285
7.3562 | -86.6276
7.3566 | -86.6283
7.3569 | -86.6292
7.3572 | -86.63
7.3575 | -86.6309
7.3579 | -86.6323
7.3582 | -86.6335
7.3585 | -86.6342
7.3589 | -86.6353
7.3592 | -86.6363
7.3595 | -86.6377
7.3599 | -86.6382
7.3602 | -86.6388
7.3605 | -86.6399
7.3608 | -86.6409
7.3612 | -86.6417
7.3615 | -86.6423
7.3618 | -86.6431
7.3622 | -86.644
7.3625 | -86.6452
7.3628 | -86.646
7.3631 | -86.6465
7.3635 | -86.6469
7.3638 | -86.6476
7.3641 | -86.6486
7.3645 | -86.6499
7.3648 | -86.6514
7.3651 | -86.6517
7.3654 | -86.6519
7.3658 | -86.6525
7.3661 | -86.6531
7.3664 | -86.6542
7.3668 | -86.6551
7.3671 | -86.6562
7.3674 | -86.6576
7.3678 | -86.6586
7.3681 | -86.6595
7.3684 | -86.6608
7.3687 | -86.662
7.3691 | -86.6634
7.3694 | -86.6643
7.3697 | -86.6652
7.3701 | -86.6658
7.3704 | -86.6665
7.3707 | -86.6676
7.371 | -86.6683
7.3714 | -86.6691
7.3717 | -86.6699
7.372 | -86.671
7.3724 | -86.6719
7.3727 | -86.6726
7.373 | -86.673
7.3733 | -86.6735
7.3737 | -86.6743
7.374 | -86.6752
7.3743 | -86.6769
7.3747 | -86.6785
7.375 | -86.6795
7.3753 | -86.6803
7.3757 | -86.6828
7.376 | -86.6864
7.3763 | -86.6885
7.3766 | -86.6902
7.377 | -86.6911
7.3773 | -86.692
7.3776 | -86.6924
7.378 | -86.6925
7.3783 | -86.693
7.3786 | -86.6944
7.3789 | -86.6952
7.3793 | -86.696
7.3796 | -86.6969
7.3799 | -86.698
7.3803 | -86.6992
7.3806 | -86.7001
7.3809 | -86.7008
7.3812 | -86.701
7.3816 | -86.7018
7.3819 | -86.7033
7.3822 | -86.7046
7.3826 | -86.7061
7.3829 | -86.7064
7.3832 | -86.7066
7.3836 | -86.7075
7.3839 | -86.7089
7.3842 | -86.7098
7.3845 | -86.7107
7.3849 | -86.7115
7.3852 | -86.7127
7.3855 | -86.7141
7.3859 | -86.7152
7.3862 | -86.716
7.3865 | -86.717
7.3868 | -86.7182
7.3872 | -86.719
7.3875 | -86.7192
7.3878 | -86.7201
7.3882 | -86.7199
7.3885 | -86.7201
7.3888 | -86.721
7.3891 | -86.7219
7.3895 | -86.7228
7.3898 | -86.7238
7.3901 | -86.7246
7.3905 | -86.7251
7.3908 | -86.7256
7.3911 | -86.7266
7.3914 | -86.7279
7.3918 | -86.7291
7.3921 | -86.7301
7.3924 | -86.731
7.3928 | -86.7325
7.3931 | -86.7335
7.3934 | -86.7342
7.3938 | -86.7356
7.3941 | -86.7362
7.3944 | -86.7369
7.3947 | -86.7383
7.3951 | -86.7383
7.3954 | -86.7397
7.3957 | -86.7411
7.3961 | -86.7416
7.3964 | -86.7435
7.3967 | -86.7444
7.397 | -86.7455
7.3974 | -86.7467
7.3977 | -86.748
7.398 | -86.7491
7.3984 | -86.75
7.3987 | -86.7507
7.399 | -86.7512
7.3994 | -86.7519
7.3997 | -86.7526
7.4 | -86.753
7.4003 | -86.7539
7.4007 | -86.7547
7.401 | -86.7552
7.4013 | -86.7553
7.4017 | -86.7564
7.402 | -86.7588
7.4023 | -86.7601
7.4026 | -86.7615
7.403 | -86.7633
7.4033 | -86.7648
7.4036 | -86.7657
7.404 | -86.7674
7.4043 | -86.7688
7.4046 | -86.77
7.4049 | -86.7708
7.4053 | -86.7715
7.4056 | -86.7723
7.4059 | -86.7732
7.4063 | -86.7744
7.4066 | -86.7752
7.4069 | -86.776
7.4072 | -86.7769
7.4076 | -86.7781
7.4079 | -86.7792
7.4082 | -86.7801
7.4086 | -86.7808
7.4089 | -86.7811
7.4092 | -86.7816
7.4096 | -86.7822
7.4099 | -86.7835
7.4102 | -86.785
7.4105 | -86.7855
7.4109 | -86.7871
7.4112 | -86.788
7.4115 | -86.7884
7.4119 | -86.7891
7.4122 | -86.7901
7.4125 | -86.7919
7.4128 | -86.7922
7.4132 | -86.7927
7.4135 | -86.7937
7.4138 | -86.7948
7.4142 | -86.7956
7.4145 | -86.7962
7.4148 | -86.7967
7.4151 | -86.7988
7.4155 | -86.8004
7.4158 | -86.8013
7.4161 | -86.8014
7.4165 | -86.8017
7.4168 | -86.802
7.4171 | -86.8033
7.4175 | -86.8052
7.4178 | -86.8061
7.4181 | -86.8064
7.4184 | -86.8072
7.4188 | -86.808
7.4191 | -86.8088
7.4194 | -86.8096
7.4198 | -86.8104
7.4201 | -86.8116
7.4204 | -86.8135
7.4207 | -86.8141
7.4211 | -86.8145
7.4214 | -86.8151
7.4217 | -86.8159
7.4221 | -86.8172
7.4224 | -86.8197
7.4227 | -86.8206
7.423 | -86.8207
7.4234 | -86.8211
7.4237 | -86.8217
7.424 | -86.8224
7.4244 | -86.8232
7.4247 | -86.8237
7.425 | -86.8241
7.4254 | -86.8249
7.4257 | -86.8259
7.426 | -86.827
7.4263 | -86.8276
7.4267 | -86.828
7.427 | -86.8289
7.4273 | -86.8298
7.4277 | -86.8306
7.428 | -86.8313
7.4283 | -86.8324
7.4286 | -86.8334
7.429 | -86.8345
7.4293 | -86.8356
7.4296 | -86.8371
7.43 | -86.8379
7.4303 | -86.8387
7.4306 | -86.8395
7.4309 | -86.8402
7.4313 | -86.8409
7.4316 | -86.8421
7.4319 | -86.8429
7.4323 | -86.8445
7.4326 | -86.8455
7.4329 | -86.8465
7.4333 | -86.8469
7.4336 | -86.8475
7.4339 | -86.8482
7.4342 | -86.8487
7.4346 | -86.8495
7.4349 | -86.8503
7.4352 | -86.8512
7.4356 | -86.8521
7.4359 | -86.8537
7.4362 | -86.8552
7.4365 | -86.8561
7.4369 | -86.8576
7.4372 | -86.8578
7.4375 | -86.8582
7.4379 | -86.8592
7.4382 | -86.8608
7.4385 | -86.8624
7.4388 | -86.8635
7.4392 | -86.8634
7.4395 | -86.8636
7.4398 | -86.8643
7.4402 | -86.8657
7.4405 | -86.8666
7.4408 | -86.8673
7.4412 | -86.8676
7.4415 | -86.8683
7.4418 | -86.8695
7.4421 | -86.8703
7.4425 | -86.8708
7.4428 | -86.8715
7.4431 | -86.8723
7.4435 | -86.8731
7.4438 | -86.8738
7.4441 | -86.8744
7.4444 | -86.8753
7.4448 | -86.8764
7.4451 | -86.8764
7.4454 | -86.8773
7.4458 | -86.8782
7.4461 | -86.8793
7.4464 | -86.8799
7.4467 | -86.8808
7.4471 | -86.8816
7.4474 | -86.8826
7.4477 | -86.8829
7.4481 | -86.8829
7.4484 | -86.8837
7.4487 | -86.8846
7.449 | -86.8856
7.4494 | -86.887
7.4497 | -86.8877
7.45 | -86.8888
7.4504 | -86.8892
7.4507 | -86.8904
7.451 | -86.893
7.4514 | -86.8947
7.4517 | -86.8954
7.452 | -86.8951
7.4523 | -86.8942
7.4527 | -86.8938
7.453 | -86.894
7.4533 | -86.8947
7.4537 | -86.8952
7.454 | -86.8959
7.4543 | -86.8967
7.4546 | -86.8979
7.455 | -86.8989
7.4553 | -86.8996
7.4556 | -86.8991
7.456 | -86.8999
7.4563 | -86.9005
7.4566 | -86.9013
7.4569 | -86.902
7.4573 | -86.9033
7.4576 | -86.906
7.4579 | -86.9078
7.4583 | -86.9088
7.4586 | -86.9095
7.4589 | -86.9104
7.4593 | -86.9114
7.4596 | -86.9124
7.4599 | -86.9131
7.4602 | -86.9146
7.4606 | -86.9155
7.4609 | -86.9161
7.4612 | -86.9172
7.4616 | -86.9183
7.4619 | -86.9196
7.4622 | -86.9208
7.4625 | -86.9216
7.4629 | -86.9224
7.4632 | -86.9237
7.4635 | -86.9245
7.4639 | -86.9253
7.4642 | -86.9254
7.4645 | -86.9265
7.4648 | -86.9277
7.4652 | -86.9283
7.4655 | -86.9292
7.4658 | -86.9304
7.4662 | -86.9309
7.4665 | -86.9317
7.4668 | -86.9323
7.4672 | -86.9332
7.4675 | -86.9344
7.4678 | -86.9353
7.4681 | -86.9358
7.4685 | -86.9363
7.4688 | -86.9361
7.4691 | -86.9364
7.4695 | -86.9367
7.4698 | -86.9376
7.4701 | -86.9392
7.4704 | -86.9403
7.4708 | -86.9411
7.4711 | -86.9429
7.4714 | -86.944
7.4718 | -86.9444
7.4721 | -86.9442
7.4724 | -86.9448
7.4727 | -86.9454
7.4731 | -86.9471
7.4734 | -86.948
7.4737 | -86.9491
7.4741 | -86.9497
7.4744 | -86.9501
7.4747 | -86.9498
7.4751 | -86.9503
7.4754 | -86.9508
7.4757 | -86.9509
7.476 | -86.9512
7.4764 | -86.952
7.4767 | -86.9513
7.477 | -86.9507
7.4774 | -86.9515
7.4777 | -86.9524
7.478 | -86.9527
7.4783 | -86.9534
7.4787 | -86.9546
7.479 | -86.9554
7.4793 | -86.956
7.4797 | -86.9567
7.48 | -86.9574
7.4803 | -86.9582
7.4806 | -86.9586
7.481 | -86.9592
7.4813 | -86.9602
7.4816 | -86.961
7.482 | -86.9616
7.4823 | -86.9624
7.4826 | -86.9631
7.483 | -86.9642
7.4833 | -86.9649
7.4836 | -86.9654
7.4839 | -86.9654
7.4843 | -86.9659
7.4846 | -86.9668
7.4849 | -86.9677
7.4853 | -86.9684
7.4856 | -86.9696
7.4859 | -86.9711
7.4862 | -86.9722
7.4866 | -86.9729
7.4869 | -86.9737
7.4872 | -86.9748
7.4876 | -86.976
7.4879 | -86.9771
7.4882 | -86.9778
7.4885 | -86.9782
7.4889 | -86.9786
7.4892 | -86.9791
7.4895 | -86.9795
7.4899 | -86.9795
7.4902 | -86.9804
7.4905 | -86.981
7.4909 | -86.9814
7.4912 | -86.9819
7.4915 | -86.9812
7.4918 | -86.9816
7.4922 | -86.982
7.4925 | -86.9828
7.4928 | -86.9833
7.4932 | -86.9841
7.4935 | -86.9846
7.4938 | -86.985
7.4941 | -86.986
7.4945 | -86.9865
7.4948 | -86.9872
7.4951 | -86.988
7.4955 | -86.9886
7.4958 | -86.9884
7.4961 | -86.9883
7.4964 | -86.9891
7.4968 | -86.9906
7.4971 | -86.9928
7.4974 | -86.9945
7.4978 | -86.9956
7.4981 | -86.9968
7.4984 | -86.9975
7.4987 | -86.9981
7.4991 | -86.9988
7.4994 | -86.9994
7.4997 | -86.9992
7.5001 | -86.999
7.5004 | -86.9989
7.5007 | -86.9989
7.5011 | -86.9994
7.5014 | -87.0003
7.5017 | -87.0008
7.502 | -87.0017
7.5024 | -87.0028
7.5027 | -87.0047
7.503 | -87.0057
7.5034 | -87.0063
7.5037 | -87.0072
7.504 | -87.0079
7.5043 | -87.0085
7.5047 | -87.0092
7.505 | -87.0096
7.5053 | -87.0104
7.5057 | -87.011
7.506 | -87.0115
7.5063 | -87.0117
7.5066 | -87.0122
7.507 | -87.0125
7.5073 | -87.0129
7.5076 | -87.0138
7.508 | -87.0139
7.5083 | -87.0146
7.5086 | -87.0151
7.509 | -87.0165
7.5093 | -87.0169
7.5096 | -87.0162
7.5099 | -87.0164
7.5103 | -87.0171
7.5106 | -87.0178
7.5109 | -87.0189
7.5113 | -87.0195
7.5116 | -87.0205
7.5119 | -87.0211
7.5122 | -87.0219
7.5126 | -87.0243
7.5129 | -87.0251
7.5132 | -87.025
7.5136 | -87.0255
7.5139 | -87.0277
7.5142 | -87.0286
7.5145 | -87.0291
7.5149 | -87.0301
7.5152 | -87.0308
7.5155 | -87.031
7.5159 | -87.0309
7.5162 | -87.0312
7.5165 | -87.0319
7.5169 | -87.0324
7.5172 | -87.0325
7.5175 | -87.0339
7.5178 | -87.0355
7.5182 | -87.036
7.5185 | -87.0355
7.5188 | -87.0353
7.5192 | -87.0357
7.5195 | -87.037
7.5198 | -87.0383
7.5201 | -87.0392
7.5205 | -87.0398
7.5208 | -87.0403
7.5211 | -87.0413
7.5215 | -87.0424
7.5218 | -87.0427
7.5221 | -87.0435
7.5224 | -87.0441
7.5228 | -87.0443
7.5231 | -87.0445
7.5234 | -87.0451
7.5238 | -87.0458
7.5241 | -87.0463
7.5244 | -87.0466
7.5248 | -87.0471
7.5251 | -87.0476
7.5254 | -87.048
7.5257 | -87.0483
7.5261 | -87.049
7.5264 | -87.0496
7.5267 | -87.0502
7.5271 | -87.0508
7.5274 | -87.0507
7.5277 | -87.0513
7.528 | -87.0516
7.5284 | -87.052
7.5287 | -87.0528
7.529 | -87.0532
7.5294 | -87.0533
7.5297 | -87.0538
7.53 | -87.0542
7.5303 | -87.0548
7.5307 | -87.055
7.531 | -87.0553
7.5313 | -87.056
7.5317 | -87.0577
7.532 | -87.0581
7.5323 | -87.0585
7.5327 | -87.0591
7.533 | -87.0601
7.5333 | -87.0607
7.5336 | -87.0613
7.534 | -87.0617
7.5343 | -87.0627
7.5346 | -87.0635
7.535 | -87.0642
7.5353 | -87.0648
7.5356 | -87.0654
7.5359 | -87.0661
7.5363 | -87.067
7.5366 | -87.0675
7.5369 | -87.0682
7.5373 | -87.0684
7.5376 | -87.069
7.5379 | -87.0694
7.5382 | -87.07
7.5386 | -87.0709
7.5389 | -87.0722
7.5392 | -87.0733
7.5396 | -87.0738
7.5399 | -87.0734
7.5402 | -87.076
7.5405 | -87.0782
7.5409 | -87.0787
7.5412 | -87.0792
7.5415 | -87.0799
7.5419 | -87.0805
7.5422 | -87.0803
7.5425 | -87.0805
7.5429 | -87.0811
7.5432 | -87.0816
7.5435 | -87.0821
7.5438 | -87.0821
7.5442 | -87.0823
7.5445 | -87.0827
7.5448 | -87.0829
7.5452 | -87.0831
7.5455 | -87.0837
7.5458 | -87.0845
7.5461 | -87.0848
7.5465 | -87.0853
7.5468 | -87.0863
7.5471 | -87.0876
7.5475 | -87.0881
7.5478 | -87.0888
7.5481 | -87.0896
7.5484 | -87.0898
7.5488 | -87.0899
7.5491 | -87.0899
7.5494 | -87.0906
7.5498 | -87.0909
7.5501 | -87.0913
7.5504 | -87.0914
7.5508 | -87.0929
7.5511 | -87.0941
7.5514 | -87.0948
7.5517 | -87.0956
7.5521 | -87.0961
7.5524 | -87.0984
7.5527 | -87.0994
7.5531 | -87.1003
7.5534 | -87.1024
7.5537 | -87.1037
7.554 | -87.1044
7.5544 | -87.105
7.5547 | -87.1057
7.555 | -87.1063
7.5554 | -87.1062
7.5557 | -87.1061
7.556 | -87.106
7.5563 | -87.1063
7.5567 | -87.1072
7.557 | -87.108
7.5573 | -87.1086
7.5577 | -87.1096
7.558 | -87.1098
7.5583 | -87.1103
7.5587 | -87.1109
7.559 | -87.1113
7.5593 | -87.1117
7.5596 | -87.1121
7.56 | -87.1126
7.5603 | -87.1128
7.5606 | -87.1132
7.561 | -87.1145
7.5613 | -87.1161
7.5616 | -87.1172
7.5619 | -87.1178
7.5623 | -87.1188
7.5626 | -87.1201
7.5629 | -87.1208
7.5633 | -87.1215
7.5636 | -87.1225
7.5639 | -87.1242
7.5642 | -87.1266
7.5646 | -87.1271
7.5649 | -87.1285
7.5652 | -87.1296
7.5656 | -87.1301
7.5659 | -87.1304
7.5662 | -87.131
7.5666 | -87.1315
7.5669 | -87.1322
7.5672 | -87.133
7.5675 | -87.1337
7.5679 | -87.1341
7.5682 | -87.1345
7.5685 | -87.1356
7.5689 | -87.1362
7.5692 | -87.1365
7.5695 | -87.1382
7.5698 | -87.14
7.5702 | -87.1407
7.5705 | -87.1412
7.5708 | -87.1417
7.5712 | -87.1423
7.5715 | -87.1429
7.5718 | -87.1433
7.5721 | -87.1436
7.5725 | -87.144
7.5728 | -87.144
7.5731 | -87.1449
7.5735 | -87.1459
7.5738 | -87.1464
7.5741 | -87.1472
7.5745 | -87.1484
7.5748 | -87.149
7.5751 | -87.1492
7.5754 | -87.149
7.5758 | -87.1499
7.5761 | -87.1515
7.5764 | -87.1519
7.5768 | -87.1529
7.5771 | -87.1537
7.5774 | -87.1544
7.5777 | -87.1553
7.5781 | -87.1565
7.5784 | -87.157
7.5787 | -87.1572
7.5791 | -87.1572
7.5794 | -87.1584
7.5797 | -87.1591
7.58 | -87.1598
7.5804 | -87.1602
7.5807 | -87.1596
7.581 | -87.1596
7.5814 | -87.1596
7.5817 | -87.16
7.582 | -87.16
7.5823 | -87.1607
7.5827 | -87.1612
7.583 | -87.1615
7.5833 | -87.1621
7.5837 | -87.1627
7.584 | -87.1634
7.5843 | -87.1642
7.5847 | -87.165
7.585 | -87.1655
7.5853 | -87.1659
7.5856 | -87.1658
7.586 | -87.1665
7.5863 | -87.1665
7.5866 | -87.1668
7.587 | -87.1681
7.5873 | -87.1691
7.5876 | -87.17
7.5879 | -87.1706
7.5883 | -87.1715
7.5886 | -87.1722
7.5889 | -87.1728
7.5893 | -87.1736
7.5896 | -87.1743
7.5899 | -87.1754
7.5903 | -87.1768
7.5906 | -87.1777
7.5909 | -87.1788
7.5912 | -87.1806
7.5916 | -87.1819
7.5919 | -87.1819
7.5922 | -87.1827
7.5926 | -87.1841
7.5929 | -87.1854
7.5932 | -87.186
7.5935 | -87.1865
7.5939 | -87.1868
7.5942 | -87.187
7.5945 | -87.1873
7.5949 | -87.1877
7.5952 | -87.1884
7.5955 | -87.1887
7.5958 | -87.1889
7.5962 | -87.1885
7.5965 | -87.1879
7.5968 | -87.1887
7.5972 | -87.1895
7.5975 | -87.1903
7.5978 | -87.1905
7.5981 | -87.1906
7.5985 | -87.1909
7.5988 | -87.1915
7.5991 | -87.192
7.5995 | -87.1928
7.5998 | -87.1933
7.6001 | -87.1943
7.6005 | -87.1953
7.6008 | -87.1961
7.6011 | -87.1969
7.6014 | -87.1969
7.6018 | -87.1973
7.6021 | -87.1981
7.6024 | -87.1988
7.6028 | -87.1992
7.6031 | -87.2002
7.6034 | -87.2009
7.6037 | -87.2016
7.6041 | -87.2025
7.6044 | -87.203
7.6047 | -87.2033
7.6051 | -87.2038
7.6054 | -87.2047
7.6057 | -87.2058
7.606 | -87.2063
7.6064 | -87.2068
7.6067 | -87.2079
7.607 | -87.209
7.6074 | -87.2101
7.6077 | -87.2116
7.608 | -87.2125
7.6084 | -87.2129
7.6087 | -87.2137
7.609 | -87.2148
7.6093 | -87.2159
7.6097 | -87.2167
7.61 | -87.2185
7.6103 | -87.2195
7.6107 | -87.2199
7.611 | -87.2204
7.6113 | -87.2208
7.6116 | -87.2208
7.612 | -87.2209
7.6123 | -87.2215
7.6126 | -87.2217
7.613 | -87.2221
7.6133 | -87.2228
7.6136 | -87.2234
7.6139 | -87.2241
7.6143 | -87.2243
7.6146 | -87.2245
7.6149 | -87.225
7.6153 | -87.2258
7.6156 | -87.2262
7.6159 | -87.2271
7.6163 | -87.2277
7.6166 | -87.2284
7.6169 | -87.2293
7.6172 | -87.23
7.6176 | -87.2307
7.6179 | -87.2308
7.6182 | -87.2308
7.6186 | -87.2311
7.6189 | -87.2318
7.6192 | -87.2323
7.6195 | -87.2328
7.6199 | -87.2332
7.6202 | -87.2335
7.6205 | -87.2347
7.6209 | -87.2357
7.6212 | -87.2363
7.6215 | -87.2373
7.6218 | -87.2381
7.6222 | -87.2383
7.6225 | -87.239
7.6228 | -87.2396
7.6232 | -87.2401
7.6235 | -87.2408
7.6238 | -87.2414
7.6242 | -87.2428
7.6245 | -87.2436
7.6248 | -87.2446
7.6251 | -87.2464
7.6255 | -87.2473
7.6258 | -87.2481
7.6261 | -87.2482
7.6265 | -87.249
7.6268 | -87.2494
7.6271 | -87.2503
7.6274 | -87.2515
7.6278 | -87.2521
7.6281 | -87.2526
7.6284 | -87.2532
7.6288 | -87.2536
7.6291 | -87.2538
7.6294 | -87.2541
7.6297 | -87.2552
7.6301 | -87.2563
7.6304 | -87.2568
7.6307 | -87.2575
7.6311 | -87.2582
7.6314 | -87.2588
7.6317 | -87.2595
7.632 | -87.2596
7.6324 | -87.26
7.6327 | -87.2612
7.633 | -87.2629
7.6334 | -87.2637
7.6337 | -87.2641
7.634 | -87.2644
7.6344 | -87.265
7.6347 | -87.2655
7.635 | -87.2651
7.6353 | -87.2657
7.6357 | -87.2664
7.636 | -87.2668
7.6363 | -87.2673
7.6367 | -87.267
7.637 | -87.2666
7.6373 | -87.2671
7.6376 | -87.2676
7.638 | -87.2676
7.6383 | -87.2681
7.6386 | -87.2692
7.639 | -87.2702
7.6393 | -87.2702
7.6396 | -87.2705
7.6399 | -87.2709
7.6403 | -87.2711
7.6406 | -87.2714
7.6409 | -87.2721
7.6413 | -87.2729
7.6416 | -87.2735
7.6419 | -87.2743
7.6423 | -87.2753
7.6426 | -87.276
7.6429 | -87.2764
7.6432 | -87.2766
7.6436 | -87.2772
7.6439 | -87.2779
7.6442 | -87.2782
7.6446 | -87.279
7.6449 | -87.2794
7.6452 | -87.28
7.6455 | -87.2807
7.6459 | -87.2813
7.6462 | -87.2819
7.6465 | -87.2825
7.6469 | -87.2828
7.6472 | -87.283
7.6475 | -87.283
7.6478 | -87.2833
7.6482 | -87.2837
7.6485 | -87.2843
7.6488 | -87.2846
7.6492 | -87.2854
7.6495 | -87.2866
7.6498 | -87.2874
7.6502 | -87.2882
7.6505 | -87.289
7.6508 | -87.2896
7.6511 | -87.2904
7.6515 | -87.2909
7.6518 | -87.2918
7.6521 | -87.2929
7.6525 | -87.2937
7.6528 | -87.294
7.6531 | -87.2946
7.6534 | -87.2953
7.6538 | -87.2958
7.6541 | -87.2955
7.6544 | -87.2958
7.6548 | -87.2964
7.6551 | -87.2971
7.6554 | -87.2974
7.6557 | -87.2975
7.6561 | -87.299
7.6564 | -87.3004
7.6567 | -87.301
7.6571 | -87.3018
7.6574 | -87.3025
7.6577 | -87.3028
7.6581 | -87.3033
7.6584 | -87.3039
7.6587 | -87.3043
7.659 | -87.3049
7.6594 | -87.3055
7.6597 | -87.306
7.66 | -87.3066
7.6604 | -87.3072
7.6607 | -87.3077
7.661 | -87.3083
7.6613 | -87.3082
7.6617 | -87.308
7.662 | -87.3086
7.6623 | -87.3096
7.6627 | -87.3102
7.663 | -87.3104
7.6633 | -87.3109
7.6636 | -87.3115
7.664 | -87.3121
7.6643 | -87.3127
7.6646 | -87.3126
7.665 | -87.3128
7.6653 | -87.314
7.6656 | -87.3146
7.666 | -87.3151
7.6663 | -87.3156
7.6666 | -87.3163
7.6669 | -87.3169
7.6673 | -87.318
7.6676 | -87.3194
7.6679 | -87.3204
7.6683 | -87.3203
7.6686 | -87.3196
7.6689 | -87.32
7.6692 | -87.3206
7.6696 | -87.3211
7.6699 | -87.3217
7.6702 | -87.3224
7.6706 | -87.3227
7.6709 | -87.3235
7.6712 | -87.3242
7.6715 | -87.325
7.6719 | -87.3258
7.6722 | -87.326
7.6725 | -87.3266
7.6729 | -87.3276
7.6732 | -87.3283
7.6735 | -87.3295
7.6739 | -87.3305
7.6742 | -87.331
7.6745 | -87.3313
7.6748 | -87.3311
7.6752 | -87.3316
7.6755 | -87.3326
7.6758 | -87.3337
7.6762 | -87.3347
7.6765 | -87.336
7.6768 | -87.3366
7.6771 | -87.3373
7.6775 | -87.3378
7.6778 | -87.3376
7.6781 | -87.3377
7.6785 | -87.3384
7.6788 | -87.339
7.6791 | -87.3394
7.6794 | -87.3402
7.6798 | -87.3409
7.6801 | -87.3414
7.6804 | -87.3421
7.6808 | -87.3427
7.6811 | -87.3429
7.6814 | -87.3439
7.6818 | -87.345
7.6821 | -87.3459
7.6824 | -87.347
7.6827 | -87.3477
7.6831 | -87.3489
7.6834 | -87.3499
7.6837 | -87.3509
7.6841 | -87.3513
7.6844 | -87.352
7.6847 | -87.3525
7.685 | -87.3531
7.6854 | -87.3534
7.6857 | -87.3544
7.686 | -87.3547
7.6864 | -87.3551
7.6867 | -87.3558
7.687 | -87.3564
7.6873 | -87.3575
7.6877 | -87.3587
7.688 | -87.3593
7.6883 | -87.3602
7.6887 | -87.3619
7.689 | -87.363
7.6893 | -87.3639
7.6896 | -87.3641
7.69 | -87.3647
7.6903 | -87.3653
7.6906 | -87.3659
7.691 | -87.3665
7.6913 | -87.3677
7.6916 | -87.3689
7.692 | -87.3695
7.6923 | -87.37
7.6926 | -87.3705
7.6929 | -87.3717
7.6933 | -87.3729
7.6936 | -87.3741
7.6939 | -87.3752
7.6943 | -87.3757
7.6946 | -87.3764
7.6949 | -87.3771
7.6952 | -87.3777
7.6956 | -87.3781
7.6959 | -87.3783
7.6962 | -87.3784
7.6966 | -87.379
7.6969 | -87.38
7.6972 | -87.3806
7.6975 | -87.3813
7.6979 | -87.3818
7.6982 | -87.3826
7.6985 | -87.3832
7.6989 | -87.3841
7.6992 | -87.3848
7.6995 | -87.3854
7.6999 | -87.3861
7.7002 | -87.387
7.7005 | -87.3883
7.7008 | -87.3891
7.7012 | -87.3896
7.7015 | -87.3908
7.7018 | -87.3926
7.7022 | -87.3946
7.7025 | -87.3955
7.7028 | -87.3964
7.7031 | -87.3977
7.7035 | -87.3983
7.7038 | -87.3986
7.7041 | -87.3992
7.7045 | -87.4001
7.7048 | -87.4009
7.7051 | -87.4028
7.7054 | -87.404
7.7058 | -87.4047
7.7061 | -87.4052
7.7064 | -87.4067
7.7068 | -87.4077
7.7071 | -87.4084
7.7074 | -87.4089
7.7078 | -87.4101
7.7081 | -87.4107
7.7084 | -87.411
7.7087 | -87.4119
7.7091 | -87.4128
7.7094 | -87.4141
7.7097 | -87.4149
7.7101 | -87.4155
7.7104 | -87.4161
7.7107 | -87.417
7.711 | -87.4179
7.7114 | -87.4186
7.7117 | -87.4194
7.712 | -87.4199
7.7124 | -87.4203
7.7127 | -87.4211
7.713 | -87.4214
7.7133 | -87.4219
7.7137 | -87.4227
7.714 | -87.4237
7.7143 | -87.4244
7.7147 | -87.4244
7.715 | -87.4251
7.7153 | -87.4258
7.7157 | -87.4261
7.716 | -87.4273
7.7163 | -87.4285
7.7166 | -87.4296
7.717 | -87.4306
7.7173 | -87.4314
7.7176 | -87.4322
7.718 | -87.4341
7.7183 | -87.4359
7.7186 | -87.4368
7.7189 | -87.4371
7.7193 | -87.4367
7.7196 | -87.4367
7.7199 | -87.4378
7.7203 | -87.4383
7.7206 | -87.4395
7.7209 | -87.4401
7.7212 | -87.4405
7.7216 | -87.4409
7.7219 | -87.4415
7.7222 | -87.4421
7.7226 | -87.443
7.7229 | -87.4434
7.7232 | -87.4444
7.7236 | -87.4457
7.7239 | -87.4468
7.7242 | -87.4483
7.7245 | -87.4493
7.7249 | -87.4504
7.7252 | -87.4529
7.7255 | -87.4556
7.7259 | -87.4564
7.7262 | -87.457
7.7265 | -87.4573
7.7268 | -87.4582
7.7272 | -87.4589
7.7275 | -87.4596
7.7278 | -87.4603
7.7282 | -87.461
7.7285 | -87.4617
7.7288 | -87.4623
7.7291 | -87.4629
7.7295 | -87.4636
7.7298 | -87.4647
7.7301 | -87.4654
7.7305 | -87.4659
7.7308 | -87.4663
7.7311 | -87.467
7.7314 | -87.468
7.7318 | -87.4691
7.7321 | -87.47
7.7324 | -87.4706
7.7328 | -87.4709
7.7331 | -87.4718
7.7334 | -87.4724
7.7338 | -87.473
7.7341 | -87.4739
7.7344 | -87.4746
7.7347 | -87.4756
7.7351 | -87.4763
7.7354 | -87.4764
7.7357 | -87.4772
7.7361 | -87.4779
7.7364 | -87.4786
7.7367 | -87.4793
7.737 | -87.4798
7.7374 | -87.4803
7.7377 | -87.481
7.738 | -87.4817
7.7384 | -87.4821
7.7387 | -87.4827
7.739 | -87.4835
7.7393 | -87.485
7.7397 | -87.4858
7.74 | -87.4866
7.7403 | -87.4874
7.7407 | -87.4886
7.741 | -87.4897
7.7413 | -87.4904
7.7417 | -87.4908
7.742 | -87.491
7.7423 | -87.4907
7.7426 | -87.4913
7.743 | -87.4915
7.7433 | -87.4919
7.7436 | -87.4925
7.744 | -87.4934
7.7443 | -87.4948
7.7446 | -87.4956
7.7449 | -87.4964
7.7453 | -87.4968
7.7456 | -87.4975
7.7459 | -87.4982
7.7463 | -87.4991
7.7466 | -87.5003
7.7469 | -87.5011
7.7472 | -87.5026
7.7476 | -87.5034
7.7479 | -87.5042
7.7482 | -87.5054
7.7486 | -87.5059
7.7489 | -87.507
7.7492 | -87.5084
7.7496 | -87.5092
7.7499 | -87.51
7.7502 | -87.5106
7.7505 | -87.5114
7.7509 | -87.5118
7.7512 | -87.5131
7.7515 | -87.5143
7.7519 | -87.5149
7.7522 | -87.5156
7.7525 | -87.5162
7.7528 | -87.5169
7.7532 | -87.5179
7.7535 | -87.5184
7.7538 | -87.5187
7.7542 | -87.5193
7.7545 | -87.5196
7.7548 | -87.5202
7.7551 | -87.5213
7.7555 | -87.5226
7.7558 | -87.5236
7.7561 | -87.5245
7.7565 | -87.5252
7.7568 | -87.5262
7.7571 | -87.5272
7.7575 | -87.5286
7.7578 | -87.5299
7.7581 | -87.5311
7.7584 | -87.5308
7.7588 | -87.5313
7.7591 | -87.5317
7.7594 | -87.5325
7.7598 | -87.5335
7.7601 | -87.5345
7.7604 | -87.5354
7.7607 | -87.5362
7.7611 | -87.5376
7.7614 | -87.5389
7.7617 | -87.5395
7.7621 | -87.5394
7.7624 | -87.5397
7.7627 | -87.5405
7.763 | -87.5419
7.7634 | -87.5429
7.7637 | -87.544
7.764 | -87.5455
7.7644 | -87.5466
7.7647 | -87.5473
7.765 | -87.5478
7.7654 | -87.5486
7.7657 | -87.549
7.766 | -87.5496
7.7663 | -87.5506
7.7667 | -87.5513
7.767 | -87.5524
7.7673 | -87.5531
7.7677 | -87.5538
7.768 | -87.5552
7.7683 | -87.556
7.7686 | -87.5574
7.769 | -87.5584
7.7693 | -87.5595
7.7696 | -87.5604
7.77 | -87.5615
7.7703 | -87.5633
7.7706 | -87.5648
7.7709 | -87.5661
7.7713 | -87.567
7.7716 | -87.5678
7.7719 | -87.5686
7.7723 | -87.5696
7.7726 | -87.5703
7.7729 | -87.5709
7.7732 | -87.5721
7.7736 | -87.5733
7.7739 | -87.5735
7.7742 | -87.5743
7.7746 | -87.5747
7.7749 | -87.5759
7.7752 | -87.5769
7.7756 | -87.5776
7.7759 | -87.5789
7.7762 | -87.5802
7.7765 | -87.581
7.7769 | -87.5814
7.7772 | -87.5822
7.7775 | -87.583
7.7779 | -87.584
7.7782 | -87.5849
7.7785 | -87.5859
7.7788 | -87.5867
7.7792 | -87.5872
7.7795 | -87.588
7.7798 | -87.5868
7.7802 | -87.5868
7.7805 | -87.5876
7.7808 | -87.5887
7.7812 | -87.5902
7.7815 | -87.5913
7.7818 | -87.5923
7.7821 | -87.5931
7.7825 | -87.5942
7.7828 | -87.596
7.7831 | -87.5976
7.7835 | -87.5986
7.7838 | -87.5993
7.7841 | -87.6004
7.7844 | -87.6012
7.7848 | -87.6024
7.7851 | -87.6034
7.7854 | -87.6033
7.7858 | -87.603
7.7861 | -87.6035
7.7864 | -87.604
7.7867 | -87.6049
7.7871 | -87.6063
7.7874 | -87.6075
7.7877 | -87.6098
7.7881 | -87.61
7.7884 | -87.6103
7.7887 | -87.6112
7.789 | -87.6122
7.7894 | -87.6133
7.7897 | -87.6142
7.79 | -87.6147
7.7904 | -87.6149
7.7907 | -87.6156
7.791 | -87.6164
7.7914 | -87.6171
7.7917 | -87.6179
7.792 | -87.6186
7.7923 | -87.6194
7.7927 | -87.6203
7.793 | -87.6215
7.7933 | -87.6225
7.7937 | -87.6234
7.794 | -87.624
7.7943 | -87.6249
7.7946 | -87.6256
7.795 | -87.6264
7.7953 | -87.6282
7.7956 | -87.6294
7.796 | -87.6306
7.7963 | -87.6316
7.7966 | -87.633
7.7969 | -87.6342
7.7973 | -87.6352
7.7976 | -87.6362
7.7979 | -87.6369
7.7983 | -87.6377
7.7986 | -87.6387
7.7989 | -87.6392
7.7993 | -87.6393
7.7996 | -87.6404
7.7999 | -87.6412
7.8002 | -87.6424
7.8006 | -87.644
7.8009 | -87.6449
7.8012 | -87.6459
7.8016 | -87.6468
7.8019 | -87.6476
7.8022 | -87.6482
7.8025 | -87.6487
7.8029 | -87.6497
7.8032 | -87.6507
7.8035 | -87.6522
7.8039 | -87.6546
7.8042 | -87.656
7.8045 | -87.6565
7.8048 | -87.6574
7.8052 | -87.6589
7.8055 | -87.6608
7.8058 | -87.6619
7.8062 | -87.6628
7.8065 | -87.6646
7.8068 | -87.6653
7.8072 | -87.6661
7.8075 | -87.6666
7.8078 | -87.6676
7.8081 | -87.6685
7.8085 | -87.6696
7.8088 | -87.6707
7.8091 | -87.672
7.8095 | -87.6734
7.8098 | -87.6744
7.8101 | -87.6757
7.8104 | -87.6766
7.8108 | -87.6766
7.8111 | -87.6779
7.8114 | -87.679
7.8118 | -87.68
7.8121 | -87.6809
7.8124 | -87.6819
7.8127 | -87.683
7.8131 | -87.6843
7.8134 | -87.6852
7.8137 | -87.6865
7.8141 | -87.6875
7.8144 | -87.689
7.8147 | -87.6899
7.8151 | -87.6909
7.8154 | -87.6922
7.8157 | -87.6931
7.816 | -87.694
7.8164 | -87.6949
7.8167 | -87.6953
7.817 | -87.6966
7.8174 | -87.6984
7.8177 | -87.6994
7.818 | -87.7002
7.8183 | -87.7018
7.8187 | -87.7032
7.819 | -87.7041
7.8193 | -87.7057
7.8197 | -87.7077
7.82 | -87.7107
7.8203 | -87.7133
7.8206 | -87.7149
7.821 | -87.7167
7.8213 | -87.7178
7.8216 | -87.7189
7.822 | -87.7214
7.8223 | -87.7227
7.8226 | -87.7233
7.8229 | -87.7234
7.8233 | -87.7231
7.8236 | -87.7234
7.8239 | -87.7248
7.8243 | -87.7279
7.8246 | -87.7307
7.8249 | -87.7312
7.8253 | -87.7319
7.8256 | -87.7326
7.8259 | -87.7338
7.8262 | -87.7361
7.8266 | -87.7383
7.8269 | -87.7392
7.8272 | -87.7401
7.8276 | -87.7408
7.8279 | -87.7417
7.8282 | -87.7425
7.8285 | -87.7436
7.8289 | -87.7444
7.8292 | -87.7454
7.8295 | -87.7468
7.8299 | -87.7476
7.8302 | -87.7482
7.8305 | -87.7493
7.8308 | -87.7501
7.8312 | -87.7512
7.8315 | -87.7525
7.8318 | -87.7539
7.8322 | -87.7553
7.8325 | -87.7564
7.8328 | -87.7573
7.8332 | -87.7579
7.8335 | -87.7587
7.8338 | -87.7603
7.8341 | -87.7613
7.8345 | -87.7613
7.8348 | -87.7619
7.8351 | -87.7633
7.8355 | -87.7647
7.8358 | -87.7656
7.8361 | -87.7663
7.8364 | -87.7672
7.8368 | -87.7682
7.8371 | -87.7691
7.8374 | -87.7695
7.8378 | -87.77
7.8381 | -87.771
7.8384 | -87.7723
7.8387 | -87.7735
7.8391 | -87.7744
7.8394 | -87.7751
7.8397 | -87.7761
7.8401 | -87.7773
7.8404 | -87.7782
7.8407 | -87.7802
7.8411 | -87.7819
7.8414 | -87.7832
7.8417 | -87.7846
7.842 | -87.7857
7.8424 | -87.7855
7.8427 | -87.7853
7.843 | -87.7856
7.8434 | -87.7864
7.8437 | -87.7872
7.844 | -87.7884
7.8443 | -87.7893
7.8447 | -87.7903
7.845 | -87.7909
7.8453 | -87.7915
7.8457 | -87.792
7.846 | -87.7934
7.8463 | -87.7956
7.8466 | -87.7966
7.847 | -87.7975
7.8473 | -87.7981
7.8476 | -87.7985
7.848 | -87.7994
7.8483 | -87.8007
7.8486 | -87.8014
7.849 | -87.8016
7.8493 | -87.8021
7.8496 | -87.803
7.8499 | -87.8038
7.8503 | -87.8046
7.8506 | -87.8067
7.8509 | -87.8085
7.8513 | -87.8093
7.8516 | -87.8106
7.8519 | -87.8117
7.8522 | -87.8127
7.8526 | -87.8134
7.8529 | -87.8141
7.8532 | -87.8153
7.8536 | -87.8163
7.8539 | -87.8172
7.8542 | -87.8182
7.8545 | -87.8193
7.8549 | -87.8203
7.8552 | -87.8216
7.8555 | -87.8223
7.8559 | -87.8223
7.8562 | -87.8227
7.8565 | -87.8235
7.8569 | -87.8244
7.8572 | -87.8252
7.8575 | -87.8258
7.8578 | -87.8254
7.8582 | -87.8257
7.8585 | -87.8265
7.8588 | -87.8267
7.8592 | -87.8275
7.8595 | -87.8275
7.8598 | -87.828
7.8601 | -87.8289
7.8605 | -87.8298
7.8608 | -87.8308
7.8611 | -87.8315
7.8615 | -87.8323
7.8618 | -87.8333
7.8621 | -87.8342
7.8624 | -87.8353
7.8628 | -87.8361
7.8631 | -87.838
7.8634 | -87.8394
7.8638 | -87.8401
7.8641 | -87.8408
7.8644 | -87.8418
7.8648 | -87.8427
7.8651 | -87.8435
7.8654 | -87.8443
7.8657 | -87.8446
7.8661 | -87.8454
7.8664 | -87.8466
7.8667 | -87.8473
7.8671 | -87.8472
7.8674 | -87.8474
7.8677 | -87.8482
7.868 | -87.8491
7.8684 | -87.8504
7.8687 | -87.8519
7.869 | -87.8514
7.8694 | -87.853
7.8697 | -87.855
7.87 | -87.8558
7.8703 | -87.8566
7.8707 | -87.8575
7.871 | -87.8579
7.8713 | -87.858
7.8717 | -87.8586
7.872 | -87.8593
7.8723 | -87.8604
7.8727 | -87.8607
7.873 | -87.8614
7.8733 | -87.8629
7.8736 | -87.8641
7.874 | -87.8646
7.8743 | -87.8655
7.8746 | -87.8668
7.875 | -87.8683
7.8753 | -87.8697
7.8756 | -87.8704
7.8759 | -87.8714
7.8763 | -87.8727
7.8766 | -87.8739
7.8769 | -87.8747
7.8773 | -87.8755
7.8776 | -87.8761
7.8779 | -87.877
7.8782 | -87.8776
7.8786 | -87.8778
7.8789 | -87.8782
7.8792 | -87.8783
7.8796 | -87.8786
7.8799 | -87.8789
7.8802 | -87.8791
7.8805 | -87.8794
7.8809 | -87.8796
7.8812 | -87.8802
7.8815 | -87.8813
7.8819 | -87.8821
7.8822 | -87.8826
7.8825 | -87.8831
7.8829 | -87.8838
7.8832 | -87.8843
7.8835 | -87.8849
7.8838 | -87.8855
7.8842 | -87.8859
7.8845 | -87.8864
7.8848 | -87.8876
7.8852 | -87.8898
7.8855 | -87.8909
7.8858 | -87.8911
7.8861 | -87.8919
7.8865 | -87.8932
7.8868 | -87.8941
7.8871 | -87.8949
7.8875 | -87.8953
7.8878 | -87.8962
7.8881 | -87.8964
7.8884 | -87.8966
7.8888 | -87.8972
7.8891 | -87.8976
7.8894 | -87.8987
7.8898 | -87.8994
7.8901 | -87.8988
7.8904 | -87.8982
7.8908 | -87.8986
7.8911 | -87.8991
7.8914 | -87.9
7.8917 | -87.9006
7.8921 | -87.9013
7.8924 | -87.9021
7.8927 | -87.9031
7.8931 | -87.9039
7.8934 | -87.9036
7.8937 | -87.9036
7.894 | -87.9043
7.8944 | -87.9051
7.8947 | -87.9062
7.895 | -87.9073
7.8954 | -87.9082
7.8957 | -87.9091
7.896 | -87.91
7.8963 | -87.9106
7.8967 | -87.9113
7.897 | -87.911
7.8973 | -87.9113
7.8977 | -87.912
7.898 | -87.9124
7.8983 | -87.913
7.8987 | -87.9134
7.899 | -87.9138
7.8993 | -87.9144
7.8996 | -87.915
7.9 | -87.9158
7.9003 | -87.9168
7.9006 | -87.9174
7.901 | -87.918
7.9013 | -87.9186
7.9016 | -87.9193
7.9019 | -87.9204
7.9023 | -87.9213
7.9026 | -87.9226
7.9029 | -87.9239
7.9033 | -87.9248
7.9036 | -87.9263
7.9039 | -87.9269
7.9042 | -87.9275
7.9046 | -87.928
7.9049 | -87.9288
7.9052 | -87.9291
7.9056 | -87.9296
7.9059 | -87.9295
7.9062 | -87.9302
7.9066 | -87.931
7.9069 | -87.9314
7.9072 | -87.932
7.9075 | -87.9327
7.9079 | -87.933
7.9082 | -87.9338
7.9085 | -87.9346
7.9089 | -87.9353
7.9092 | -87.9359
7.9095 | -87.9367
7.9098 | -87.9373
7.9102 | -87.9376
7.9105 | -87.9379
7.9108 | -87.9385
7.9112 | -87.939
7.9115 | -87.9393
7.9118 | -87.9396
7.9121 | -87.941
7.9125 | -87.9415
7.9128 | -87.942
7.9131 | -87.943
7.9135 | -87.9435
7.9138 | -87.9443
7.9141 | -87.9448
7.9145 | -87.9455
7.9148 | -87.9462
7.9151 | -87.9465
7.9154 | -87.947
7.9158 | -87.9475
7.9161 | -87.9476
7.9164 | -87.9481
7.9168 | -87.9488
7.9171 | -87.9489
7.9174 | -87.9493
7.9177 | -87.9499
7.9181 | -87.9504
7.9184 | -87.9515
7.9187 | -87.9521
7.9191 | -87.9534
7.9194 | -87.9551
7.9197 | -87.9551
7.92 | -87.9568
7.9204 | -87.9576
7.9207 | -87.9581
7.921 | -87.9586
7.9214 | -87.9589
7.9217 | -87.9597
7.922 | -87.9604
7.9223 | -87.9611
7.9227 | -87.9614
7.923 | -87.9612
7.9233 | -87.9604
7.9237 | -87.9603
7.924 | -87.9609
7.9243 | -87.9617
7.9247 | -87.9626
7.925 | -87.9632
7.9253 | -87.9633
7.9256 | -87.9645
7.926 | -87.966
7.9263 | -87.9672
7.9266 | -87.9689
7.927 | -87.969
7.9273 | -87.9685
7.9276 | -87.9682
7.9279 | -87.9686
7.9283 | -87.9699
7.9286 | -87.9703
7.9289 | -87.9705
7.9293 | -87.9692
7.9296 | -87.9683
7.9299 | -87.9675
7.9302 | -87.9678
7.9306 | -87.9684
7.9309 | -87.9689
7.9312 | -87.9695
7.9316 | -87.97
7.9319 | -87.9707
7.9322 | -87.9717
7.9326 | -87.9719
7.9329 | -87.9723
7.9332 | -87.9733
7.9335 | -87.9747
7.9339 | -87.9753
7.9342 | -87.9755
7.9345 | -87.9762
7.9349 | -87.9766
7.9352 | -87.9769
7.9355 | -87.9775
7.9358 | -87.9781
7.9362 | -87.9786
7.9365 | -87.9791
7.9368 | -87.9793
7.9372 | -87.9792
7.9375 | -87.9798
7.9378 | -87.9794
7.9381 | -87.9792
7.9385 | -87.9794
7.9388 | -87.9793
7.9391 | -87.9795
7.9395 | -87.9798
7.9398 | -87.9802
7.9401 | -87.9806
7.9405 | -87.9804
7.9408 | -87.9807
7.9411 | -87.9813
7.9414 | -87.9816
7.9418 | -87.9826
7.9421 | -87.9836
7.9424 | -87.9845
7.9428 | -87.985
7.9431 | -87.9851
7.9434 | -87.9855
7.9437 | -87.9857
7.9441 | -87.9862
7.9444 | -87.9861
7.9447 | -87.986
7.9451 | -87.9863
7.9454 | -87.9866
7.9457 | -87.988
7.946 | -87.9887
7.9464 | -87.9888
7.9467 | -87.9888
7.947 | -87.9892
7.9474 | -87.9897
7.9477 | -87.9906
7.948 | -87.991
7.9484 | -87.9905
7.9487 | -87.9905
7.949 | -87.9909
7.9493 | -87.9919
7.9497 | -87.9925
7.95 | -87.9935
7.9503 | -87.9943
7.9507 | -87.9939
7.951 | -87.9934
7.9513 | -87.9938
7.9516 | -87.995
7.952 | -87.9954
7.9523 | -87.9953
7.9526 | -87.9954
7.953 | -87.9957
7.9533 | -87.9952
7.9536 | -87.9949
7.9539 | -87.9954
7.9543 | -87.9965
7.9546 | -87.9964
7.9549 | -87.9968
7.9553 | -87.9973
7.9556 | -87.9974
7.9559 | -87.9978
7.9563 | -87.9982
7.9566 | -87.9995
7.9569 | -88.0002
7.9572 | -88.0009
7.9576 | -88.0012
7.9579 | -88.0009
7.9582 | -88.0012
7.9586 | -88.0021
7.9589 | -88.0026
7.9592 | -88.0027
7.9595 | -88.0031
7.9599 | -88.003
7.9602 | -88.0034
7.9605 | -88.0041
7.9609 | -88.0049
7.9612 | -88.005
7.9615 | -88.005
7.9618 | -88.0052
7.9622 | -88.0058
7.9625 | -88.0062
7.9628 | -88.0065
7.9632 | -88.0067
7.9635 | -88.007
7.9638 | -88.0076
7.9642 | -88.008
7.9645 | -88.0082
7.9648 | -88.0082
7.9651 | -88.0083
7.9655 | -88.0086
7.9658 | -88.0086
7.9661 | -88.0086
7.9665 | -88.0092
7.9668 | -88.0097
7.9671 | -88.0097
7.9674 | -88.0101
7.9678 | -88.0108
7.9681 | -88.0112
7.9684 | -88.0109
7.9688 | -88.011
7.9691 | -88.0113
7.9694 | -88.0113
7.9697 | -88.0113
7.9701 | -88.0119
7.9704 | -88.0121
7.9707 | -88.0124
7.9711 | -88.0125
7.9714 | -88.0126
7.9717 | -88.0128
7.9721 | -88.0142
7.9724 | -88.0152
7.9727 | -88.0155
7.973 | -88.0157
7.9734 | -88.0156
7.9737 | -88.0156
7.974 | -88.0167
7.9744 | -88.0179
7.9747 | -88.0182
7.975 | -88.0183
7.9753 | -88.0185
7.9757 | -88.0189
7.976 | -88.0198
7.9763 | -88.0203
7.9767 | -88.0206
7.977 | -88.021
7.9773 | -88.0211
7.9776 | -88.0212
7.978 | -88.0213
7.9783 | -88.0213
7.9786 | -88.0214
7.979 | -88.0216
7.9793 | -88.0217
7.9796 | -88.0208
7.9799 | -88.0205
7.9803 | -88.0203
7.9806 | -88.0209
7.9809 | -88.021
7.9813 | -88.0209
7.9816 | -88.0209
7.9819 | -88.0209
7.9823 | -88.021
7.9826 | -88.0214
7.9829 | -88.0216
7.9832 | -88.0225
7.9836 | -88.024
7.9839 | -88.0241
7.9842 | -88.0239
7.9846 | -88.0239
7.9849 | -88.0237
7.9852 | -88.0222
7.9855 | -88.0209
7.9859 | -88.0203
7.9862 | -88.018
7.9865 | -88.0186
7.9869 | -88.0187
7.9872 | -88.0188
7.9875 | -88.0195
7.9878 | -88.0208
7.9882 | -88.0227
7.9885 | -88.0234
7.9888 | -88.024
7.9892 | -88.0239
7.9895 | -88.0243
7.9898 | -88.0243
7.9902 | -88.0243
7.9905 | -88.0243
7.9908 | -88.0244
7.9911 | -88.0246
7.9915 | -88.0249
7.9918 | -88.0262
7.9921 | -88.0266
7.9925 | -88.0268
7.9928 | -88.0268
7.9931 | -88.0266
7.9934 | -88.0265
7.9938 | -88.0261
7.9941 | -88.0259
7.9944 | -88.0256
7.9948 | -88.0254
7.9951 | -88.0254
7.9954 | -88.0254
7.9957 | -88.0257
7.9961 | -88.026
7.9964 | -88.0265
7.9967 | -88.0265
7.9971 | -88.0267
7.9974 | -88.027
7.9977 | -88.0271
7.9981 | -88.0271
7.9984 | -88.0273
7.9987 | -88.0272
7.999 | -88.0271
7.9994 | -88.027
7.9997 | -88.0267
8 | -88.0266
8.0004 | -88.0268
8.0007 | -88.027
8.001 | -88.0277
8.0013 | -88.0281
8.0017 | -88.028
8.002 | -88.0281
8.0023 | -88.0282
8.0027 | -88.0278
8.003 | -88.0271
8.0033 | -88.0275
8.0036 | -88.0281
8.004 | -88.0282
8.0043 | -88.0281
8.0046 | -88.0269
8.005 | -88.0257
8.0053 | -88.0252
8.0056 | -88.0259
8.006 | -88.0271
8.0063 | -88.0272
8.0066 | -88.0275
8.0069 | -88.0277
8.0073 | -88.0276
8.0076 | -88.0269
8.0079 | -88.0263
8.0083 | -88.0261
8.0086 | -88.0256
8.0089 | -88.0259
8.0092 | -88.0274
8.0096 | -88.0277
8.0099 | -88.0274
8.0102 | -88.0272
8.0106 | -88.0271
8.0109 | -88.0268
8.0112 | -88.0264
8.0115 | -88.0266
8.0119 | -88.0265
8.0122 | -88.0279
8.0125 | -88.0286
8.0129 | -88.0289
8.0132 | -88.0297
8.0135 | -88.0299
8.0138 | -88.0297
8.0142 | -88.0315
8.0145 | -88.0326
8.0148 | -88.0323
8.0152 | -88.0324
8.0155 | -88.0323
8.0158 | -88.0322
8.0162 | -88.0322
8.0165 | -88.0321
8.0168 | -88.0323
8.0171 | -88.0321
8.0175 | -88.0314
8.0178 | -88.0312
8.0181 | -88.0316
8.0185 | -88.0323
8.0188 | -88.0325
8.0191 | -88.0324
8.0194 | -88.0321
8.0198 | -88.0315
8.0201 | -88.031
8.0204 | -88.0309
8.0208 | -88.0308
8.0211 | -88.0307
8.0214 | -88.0305
8.0217 | -88.0306
8.0221 | -88.0319
8.0224 | -88.0329
8.0227 | -88.0331
8.0231 | -88.0329
8.0234 | -88.0331
8.0237 | -88.0331
8.0241 | -88.0328
8.0244 | -88.0328
8.0247 | -88.0317
8.025 | -88.031
8.0254 | -88.0305
8.0257 | -88.0302
8.026 | -88.0304
8.0264 | -88.0296
8.0267 | -88.0283
8.027 | -88.0277
8.0273 | -88.028
8.0277 | -88.0281
8.028 | -88.0287
8.0283 | -88.0297
8.0287 | -88.0296
8.029 | -88.0296
8.0293 | -88.0295
8.0296 | -88.0293
8.03 | -88.0307
8.0303 | -88.0306
8.0306 | -88.0304
8.031 | -88.0304
8.0313 | -88.0303
8.0316 | -88.0303
8.032 | -88.03
8.0323 | -88.0302
8.0326 | -88.0296
8.0329 | -88.0288
8.0333 | -88.0288
8.0336 | -88.0286
8.0339 | -88.0286
8.0343 | -88.0285
8.0346 | -88.0285
8.0349 | -88.0294
8.0352 | -88.0303
8.0356 | -88.0302
8.0359 | -88.03
8.0362 | -88.0298
8.0366 | -88.0298
8.0369 | -88.0295
8.0372 | -88.0284
8.0375 | -88.0272
8.0379 | -88.0265
8.0382 | -88.0259
8.0385 | -88.0259
8.0389 | -88.0268
8.0392 | -88.0278
8.0395 | -88.0289
8.0399 | -88.0293
8.0402 | -88.0295
8.0405 | -88.0294
8.0408 | -88.0295
8.0412 | -88.0295
8.0415 | -88.0294
8.0418 | -88.0292
8.0422 | -88.0289
8.0425 | -88.0286
8.0428 | -88.0282
8.0431 | -88.0274
8.0435 | -88.0276
8.0438 | -88.0275
8.0441 | -88.0274
8.0445 | -88.0272
8.0448 | -88.0271
8.0451 | -88.0268
8.0454 | -88.0263
8.0458 | -88.0263
8.0461 | -88.0267
8.0464 | -88.0272
8.0468 | -88.0283
8.0471 | -88.0282
8.0474 | -88.028
8.0478 | -88.0284
8.0481 | -88.028
8.0484 | -88.0275
8.0487 | -88.0266
8.0491 | -88.0271
8.0494 | -88.0287
8.0497 | -88.0288
8.0501 | -88.0281
8.0504 | -88.0277
8.0507 | -88.0278
8.051 | -88.0278
8.0514 | -88.0277
8.0517 | -88.0276
8.052 | -88.0271
8.0524 | -88.0267
8.0527 | -88.027
8.053 | -88.0271
8.0533 | -88.0275
8.0537 | -88.0282
8.054 | -88.0287
8.0543 | -88.0291
8.0547 | -88.0284
8.055 | -88.0274
8.0553 | -88.0271
8.0557 | -88.0272
8.056 | -88.0273
8.0563 | -88.0271
8.0566 | -88.0265
8.057 | -88.0263
8.0573 | -88.0266
8.0576 | -88.0265
8.058 | -88.0259
8.0583 | -88.0259
8.0586 | -88.0267
8.0589 | -88.028
8.0593 | -88.0274
8.0596 | -88.0271
8.0599 | -88.0274
8.0603 | -88.0274
8.0606 | -88.0271
8.0609 | -88.0271
8.0612 | -88.0272
8.0616 | -88.0266
8.0619 | -88.0261
8.0622 | -88.0258
8.0626 | -88.0256
8.0629 | -88.026
8.0632 | -88.0254
8.0635 | -88.0242
8.0639 | -88.0242
8.0642 | -88.0243
8.0645 | -88.0236
8.0649 | -88.0237
8.0652 | -88.0231
8.0655 | -88.0226
8.0659 | -88.0218
8.0662 | -88.0206
8.0665 | -88.02
8.0668 | -88.02
8.0672 | -88.0196
8.0675 | -88.0197
8.0678 | -88.0199
8.0682 | -88.02
8.0685 | -88.0203
8.0688 | -88.0203
8.0691 | -88.0205
8.0695 | -88.0197
8.0698 | -88.0185
8.0701 | -88.0179
8.0705 | -88.0173
8.0708 | -88.0161
8.0711 | -88.0153
8.0714 | -88.0149
8.0718 | -88.0146
8.0721 | -88.014
8.0724 | -88.014
8.0728 | -88.014
8.0731 | -88.0139
8.0734 | -88.0133
8.0738 | -88.0128
8.0741 | -88.0126
8.0744 | -88.0125
8.0747 | -88.0121
8.0751 | -88.0115
8.0754 | -88.0115
8.0757 | -88.0117
8.0761 | -88.0114
8.0764 | -88.0113
8.0767 | -88.0108
8.077 | -88.0103
8.0774 | -88.0096
8.0777 | -88.0092
8.078 | -88.0088
8.0784 | -88.0083
8.0787 | -88.0081
8.079 | -88.0081
8.0793 | -88.008
8.0797 | -88.0079
8.08 | -88.0077
8.0803 | -88.0075
8.0807 | -88.0072
8.081 | -88.007
8.0813 | -88.0056
8.0817 | -88.0052
8.082 | -88.0052
8.0823 | -88.0053
8.0826 | -88.0052
8.083 | -88.0045
8.0833 | -88.004
8.0836 | -88.0037
8.084 | -88.0035
8.0843 | -88.0035
8.0846 | -88.0041
8.0849 | -88.0041
8.0853 | -88.0038
8.0856 | -88.0038
8.0859 | -88.0035
8.0863 | -88.0037
8.0866 | -88.0039
8.0869 | -88.0042
8.0872 | -88.0046
8.0876 | -88.0038
8.0879 | -88.0027
8.0882 | -88.0025
8.0886 | -88.0025
8.0889 | -88.0025
8.0892 | -88.0026
8.0896 | -88.0022
8.0899 | -88.002
8.0902 | -88.0019
8.0905 | -88.0017
8.0909 | -88.0012
8.0912 | -88.0012
8.0915 | -88.0005
8.0919 | -88.0004
8.0922 | -88.0003
8.0925 | -87.9996
8.0928 | -87.999
8.0932 | -87.9986
8.0935 | -87.9988
8.0938 | -87.9979
8.0942 | -87.9972
8.0945 | -87.9969
8.0948 | -87.9968
8.0951 | -87.9963
8.0955 | -87.9951
8.0958 | -87.9943
8.0961 | -87.9941
8.0965 | -87.9943
8.0968 | -87.9941
8.0971 | -87.9937
8.0975 | -87.9933
8.0978 | -87.9931
8.0981 | -87.9931
8.0984 | -87.9934
8.0988 | -87.9941
8.0991 | -87.9946
8.0994 | -87.9948
8.0998 | -87.9946
8.1001 | -87.9945
8.1004 | -87.9945
8.1007 | -87.9947
8.1011 | -87.9946
8.1014 | -87.9941
8.1017 | -87.9937
8.1021 | -87.9934
8.1024 | -87.9933
8.1027 | -87.9935
8.103 | -87.9934
8.1034 | -87.9934
8.1037 | -87.9932
8.104 | -87.9933
8.1044 | -87.993
8.1047 | -87.9919
8.105 | -87.9916
8.1053 | -87.9913
8.1057 | -87.9912
8.106 | -87.9912
8.1063 | -87.991
8.1067 | -87.9909
8.107 | -87.9907
8.1073 | -87.9902
8.1077 | -87.9899
8.108 | -87.9898
8.1083 | -87.9898
8.1086 | -87.9896
8.109 | -87.9897
8.1093 | -87.9894
8.1096 | -87.9894
8.11 | -87.9892
8.1103 | -87.9888
8.1106 | -87.9887
8.1109 | -87.9883
8.1113 | -87.9873
8.1116 | -87.9865
8.1119 | -87.9865
8.1123 | -87.9864
8.1126 | -87.986
8.1129 | -87.9857
8.1133 | -87.9855
8.1136 | -87.9853
8.1139 | -87.985
8.1142 | -87.9845
8.1146 | -87.9843
8.1149 | -87.9844
8.1152 | -87.984
8.1156 | -87.9838
8.1159 | -87.9836
8.1162 | -87.9839
8.1165 | -87.9841
8.1169 | -87.9838
8.1172 | -87.9839
8.1175 | -87.9834
8.1179 | -87.9831
8.1182 | -87.9829
8.1185 | -87.9825
8.1188 | -87.9822
8.1192 | -87.9824
8.1195 | -87.9819
8.1198 | -87.9818
8.1202 | -87.9826
8.1205 | -87.9827
8.1208 | -87.9822
8.1211 | -87.9813
8.1215 | -87.981
8.1218 | -87.9809
8.1221 | -87.9807
8.1225 | -87.9801
8.1228 | -87.9795
8.1231 | -87.9791
8.1235 | -87.9788
8.1238 | -87.9787
8.1241 | -87.9791
8.1244 | -87.9789
8.1248 | -87.9782
8.1251 | -87.9777
8.1254 | -87.9778
8.1258 | -87.978
8.1261 | -87.9773
8.1264 | -87.9776
8.1267 | -87.9779
8.1271 | -87.9784
8.1274 | -87.978
8.1277 | -87.9779
8.1281 | -87.978
8.1284 | -87.9773
8.1287 | -87.977
8.129 | -87.977
8.1294 | -87.9763
8.1297 | -87.9758
8.13 | -87.9759
8.1304 | -87.9757
8.1307 | -87.9758
8.131 | -87.9755
8.1314 | -87.9749
8.1317 | -87.9746
8.132 | -87.9746
8.1323 | -87.9748
8.1327 | -87.9748
8.133 | -87.9744
8.1333 | -87.9734
8.1337 | -87.972
8.134 | -87.9704
8.1343 | -87.9693
8.1346 | -87.9688
8.135 | -87.9683
8.1353 | -87.9671
8.1356 | -87.9659
8.136 | -87.9652
8.1363 | -87.964
8.1366 | -87.9634
8.1369 | -87.9635
8.1373 | -87.9628
8.1376 | -87.9627
8.1379 | -87.9627
8.1383 | -87.9618
8.1386 | -87.962
8.1389 | -87.962
8.1393 | -87.9619
8.1396 | -87.9619
8.1399 | -87.9618
8.1402 | -87.9619
8.1406 | -87.9619
8.1409 | -87.9613
8.1412 | -87.9608
8.1416 | -87.9607
8.1419 | -87.9607
8.1422 | -87.9611
8.1425 | -87.9614
8.1429 | -87.9611
8.1432 | -87.9608
8.1435 | -87.9607
8.1439 | -87.9599
8.1442 | -87.9598
8.1445 | -87.9601
8.1448 | -87.96
8.1452 | -87.9596
8.1455 | -87.9589
8.1458 | -87.959
8.1462 | -87.9589
8.1465 | -87.9585
8.1468 | -87.9582
8.1472 | -87.9583
8.1475 | -87.9581
8.1478 | -87.9585
8.1481 | -87.9586
8.1485 | -87.958
8.1488 | -87.9574
8.1491 | -87.9569
8.1495 | -87.9568
8.1498 | -87.9565
8.1501 | -87.956
8.1504 | -87.9559
8.1508 | -87.9561
8.1511 | -87.9563
8.1514 | -87.9564
8.1518 | -87.9562
8.1521 | -87.956
8.1524 | -87.9557
8.1527 | -87.9547
8.1531 | -87.954
8.1534 | -87.9529
8.1537 | -87.9521
8.1541 | -87.9522
8.1544 | -87.9519
8.1547 | -87.9515
8.155 | -87.9513
8.1554 | -87.9513
8.1557 | -87.9509
8.156 | -87.9505
8.1564 | -87.9492
8.1567 | -87.9486
8.157 | -87.9485
8.1574 | -87.9481
8.1577 | -87.9474
8.158 | -87.947
8.1583 | -87.947
8.1587 | -87.9473
8.159 | -87.9474
8.1593 | -87.9475
8.1597 | -87.9472
8.16 | -87.9469
8.1603 | -87.9467
8.1606 | -87.9463
8.161 | -87.9461
8.1613 | -87.9462
8.1616 | -87.9463
8.162 | -87.9465
8.1623 | -87.9466
8.1626 | -87.9461
8.163 | -87.9454
8.1633 | -87.9452
8.1636 | -87.9447
8.1639 | -87.9442
8.1643 | -87.9438
8.1646 | -87.9435
8.1649 | -87.9423
8.1653 | -87.9424
8.1656 | -87.9436
8.1659 | -87.9439
8.1662 | -87.9435
8.1666 | -87.9422
8.1669 | -87.9422
8.1672 | -87.942
8.1676 | -87.9417
8.1679 | -87.9417
8.1682 | -87.9418
8.1685 | -87.9417
8.1689 | -87.9416
8.1692 | -87.9413
8.1695 | -87.9411
8.1699 | -87.9409
8.1702 | -87.941
8.1705 | -87.9406
8.1708 | -87.9404
8.1712 | -87.9404
8.1715 | -87.94
8.1718 | -87.9399
8.1722 | -87.9399
8.1725 | -87.9394
8.1728 | -87.9395
8.1732 | -87.9397
8.1735 | -87.9396
8.1738 | -87.9395
8.1741 | -87.9396
8.1745 | -87.9397
8.1748 | -87.9396
8.1751 | -87.9399
8.1755 | -87.9402
8.1758 | -87.9408
8.1761 | -87.9392
8.1764 | -87.9389
8.1768 | -87.9386
8.1771 | -87.9384
8.1774 | -87.9379
8.1778 | -87.9375
8.1781 | -87.9374
8.1784 | -87.9371
8.1787 | -87.9369
8.1791 | -87.9367
8.1794 | -87.9362
8.1797 | -87.936
8.1801 | -87.9358
8.1804 | -87.9359
8.1807 | -87.9361
8.1811 | -87.9355
8.1814 | -87.9352
8.1817 | -87.9353
8.182 | -87.9341
8.1824 | -87.9341
8.1827 | -87.9341
8.183 | -87.9342
8.1834 | -87.9343
8.1837 | -87.9346
8.184 | -87.9353
8.1843 | -87.9348
8.1847 | -87.9347
8.185 | -87.9347
8.1853 | -87.9344
8.1857 | -87.9344
8.186 | -87.9351
8.1863 | -87.9358
8.1866 | -87.9352
8.187 | -87.935
8.1873 | -87.935
8.1876 | -87.9337
8.188 | -87.9337
8.1883 | -87.9332
8.1886 | -87.933
8.189 | -87.9329
8.1893 | -87.9325
8.1896 | -87.9327
8.1899 | -87.9325
8.1903 | -87.9322
8.1906 | -87.932
8.1909 | -87.9318
8.1913 | -87.9315
8.1916 | -87.931
8.1919 | -87.9303
8.1922 | -87.93
8.1926 | -87.93
8.1929 | -87.9297
8.1932 | -87.9296
8.1936 | -87.9295
8.1939 | -87.9294
8.1942 | -87.9293
8.1945 | -87.9289
8.1949 | -87.9286
8.1952 | -87.9273
8.1955 | -87.927
8.1959 | -87.9263
8.1962 | -87.9241
8.1965 | -87.9225
8.1969 | -87.9223
8.1972 | -87.9226
8.1975 | -87.9235
8.1978 | -87.9244
8.1982 | -87.9245
8.1985 | -87.924
8.1988 | -87.9236
8.1992 | -87.9232
8.1995 | -87.9227
8.1998 | -87.9225
8.2001 | -87.9216
8.2005 | -87.9204
8.2008 | -87.9197
8.2011 | -87.9193
8.2015 | -87.919
8.2018 | -87.9186
8.2021 | -87.9184
8.2024 | -87.9184
8.2028 | -87.9183
8.2031 | -87.9177
8.2034 | -87.9172
8.2038 | -87.9168
8.2041 | -87.9168
8.2044 | -87.9176
8.2048 | -87.9172
8.2051 | -87.9174
8.2054 | -87.917
8.2057 | -87.9167
8.2061 | -87.9166
8.2064 | -87.9168
8.2067 | -87.9169
8.2071 | -87.917
8.2074 | -87.9173
8.2077 | -87.9171
8.208 | -87.9167
8.2084 | -87.9165
8.2087 | -87.9162
8.209 | -87.9159
8.2094 | -87.9157
8.2097 | -87.9154
8.21 | -87.9152
8.2103 | -87.915
8.2107 | -87.9148
8.211 | -87.9145
8.2113 | -87.9143
8.2117 | -87.9133
8.212 | -87.9127
8.2123 | -87.9132
8.2126 | -87.9129
8.213 | -87.9127
8.2133 | -87.9128
8.2136 | -87.9127
8.214 | -87.9126
8.2143 | -87.9127
8.2146 | -87.9127
8.215 | -87.9124
8.2153 | -87.9123
8.2156 | -87.9125
8.2159 | -87.9132
8.2163 | -87.9134
8.2166 | -87.9132
8.2169 | -87.9127
8.2173 | -87.9123
8.2176 | -87.9122
8.2179 | -87.9114
8.2182 | -87.911
8.2186 | -87.9108
8.2189 | -87.9106
8.2192 | -87.91
8.2196 | -87.91
8.2199 | -87.9108
8.2202 | -87.9106
8.2205 | -87.9103
8.2209 | -87.9099
8.2212 | -87.9097
8.2215 | -87.9093
8.2219 | -87.9092
8.2222 | -87.9088
8.2225 | -87.9083
8.2229 | -87.9074
8.2232 | -87.9069
8.2235 | -87.907
8.2238 | -87.9071
8.2242 | -87.9068
8.2245 | -87.9066
8.2248 | -87.9062
8.2252 | -87.9059
8.2255 | -87.9055
8.2258 | -87.905
8.2261 | -87.9045
8.2265 | -87.9034
8.2268 | -87.9039
8.2271 | -87.9042
8.2275 | -87.9041
8.2278 | -87.9038
8.2281 | -87.9034
8.2284 | -87.9026
8.2288 | -87.9019
8.2291 | -87.9014
8.2294 | -87.9013
8.2298 | -87.9017
8.2301 | -87.9011
8.2304 | -87.9005
8.2308 | -87.9001
8.2311 | -87.9
8.2314 | -87.8996
8.2317 | -87.8993
8.2321 | -87.899
8.2324 | -87.8979
8.2327 | -87.8972
8.2331 | -87.8973
8.2334 | -87.897
8.2337 | -87.8966
8.234 | -87.8965
8.2344 | -87.8958
8.2347 | -87.8943
8.235 | -87.8936
8.2354 | -87.8932
8.2357 | -87.8927
8.236 | -87.8922
8.2363 | -87.8917
8.2367 | -87.8913
8.237 | -87.8912
8.2373 | -87.8911
8.2377 | -87.8906
8.238 | -87.8908
8.2383 | -87.8913
8.2387 | -87.8916
8.239 | -87.8915
8.2393 | -87.8911
8.2396 | -87.8905
8.24 | -87.89
8.2403 | -87.8898
8.2406 | -87.8898
8.241 | -87.8897
8.2413 | -87.8894
8.2416 | -87.8891
8.2419 | -87.889
8.2423 | -87.8893
8.2426 | -87.8892
8.2429 | -87.8891
8.2433 | -87.8882
8.2436 | -87.8878
8.2439 | -87.8875
8.2442 | -87.8871
8.2446 | -87.8865
8.2449 | -87.8867
8.2452 | -87.8862
8.2456 | -87.886
8.2459 | -87.8858
8.2462 | -87.8855
8.2466 | -87.8855
8.2469 | -87.8852
8.2472 | -87.885
8.2475 | -87.8848
8.2479 | -87.8843
8.2482 | -87.8837
8.2485 | -87.8836
8.2489 | -87.8833
8.2492 | -87.8833
8.2495 | -87.8832
8.2498 | -87.883
8.2502 | -87.8826
8.2505 | -87.8825
8.2508 | -87.8825
8.2512 | -87.8822
8.2515 | -87.8814
8.2518 | -87.8808
8.2521 | -87.8808
8.2525 | -87.8805
8.2528 | -87.8805
8.2531 | -87.8802
8.2535 | -87.8799
8.2538 | -87.8798
8.2541 | -87.8795
8.2545 | -87.8788
8.2548 | -87.8783
8.2551 | -87.8781
8.2554 | -87.8777
8.2558 | -87.8771
8.2561 | -87.8765
8.2564 | -87.8762
8.2568 | -87.8761
8.2571 | -87.8757
8.2574 | -87.8754
8.2577 | -87.8755
8.2581 | -87.8752
8.2584 | -87.8748
8.2587 | -87.8744
8.2591 | -87.8741
8.2594 | -87.874
8.2597 | -87.8741
8.26 | -87.8737
8.2604 | -87.872
8.2607 | -87.8711
8.261 | -87.8705
8.2614 | -87.8702
8.2617 | -87.8702
8.262 | -87.8697
8.2623 | -87.87
8.2627 | -87.87
8.263 | -87.8695
8.2633 | -87.8691
8.2637 | -87.8684
8.264 | -87.8682
8.2643 | -87.8684
8.2647 | -87.8676
8.265 | -87.8672
8.2653 | -87.8669
8.2656 | -87.8663
8.266 | -87.8658
8.2663 | -87.865
8.2666 | -87.8642
8.267 | -87.8634
8.2673 | -87.8628
8.2676 | -87.8627
8.2679 | -87.8622
8.2683 | -87.8614
8.2686 | -87.8609
8.2689 | -87.8603
8.2693 | -87.8606
8.2696 | -87.86
8.2699 | -87.8595
8.2702 | -87.8594
8.2706 | -87.8582
8.2709 | -87.8576
8.2712 | -87.8574
8.2716 | -87.8566
8.2719 | -87.8556
8.2722 | -87.8552
8.2726 | -87.8549
8.2729 | -87.8544
8.2732 | -87.8541
8.2735 | -87.8537
8.2739 | -87.8538
8.2742 | -87.8547
8.2745 | -87.8552
8.2749 | -87.8549
8.2752 | -87.8548
8.2755 | -87.8545
8.2758 | -87.8539
8.2762 | -87.853
8.2765 | -87.8526
8.2768 | -87.8527
8.2772 | -87.8522
8.2775 | -87.852
8.2778 | -87.8518
8.2781 | -87.8515
8.2785 | -87.851
8.2788 | -87.8507
8.2791 | -87.8503
8.2795 | -87.85
8.2798 | -87.8499
8.2801 | -87.8498
8.2805 | -87.8497
8.2808 | -87.8492
8.2811 | -87.8482
8.2814 | -87.8478
8.2818 | -87.8474
8.2821 | -87.8474
8.2824 | -87.8463
8.2828 | -87.8457
8.2831 | -87.845
8.2834 | -87.8446
8.2837 | -87.8441
8.2841 | -87.8438
8.2844 | -87.8437
8.2847 | -87.8434
8.2851 | -87.8434
8.2854 | -87.8422
8.2857 | -87.8407
8.286 | -87.8399
8.2864 | -87.8389
8.2867 | -87.8386
8.287 | -87.8381
8.2874 | -87.8375
8.2877 | -87.8371
8.288 | -87.837
8.2884 | -87.8368
8.2887 | -87.8368
8.289 | -87.8363
8.2893 | -87.8357
8.2897 | -87.8361
8.29 | -87.8358
8.2903 | -87.8358
8.2907 | -87.8354
8.291 | -87.8351
8.2913 | -87.8349
8.2916 | -87.8348
8.292 | -87.8341
8.2923 | -87.8334
8.2926 | -87.8332
8.293 | -87.8332
8.2933 | -87.8334
8.2936 | -87.8334
8.2939 | -87.8331
8.2943 | -87.8328
8.2946 | -87.8328
8.2949 | -87.8326
8.2953 | -87.8322
8.2956 | -87.8313
8.2959 | -87.8305
8.2963 | -87.8304
8.2966 | -87.8299
8.2969 | -87.8294
8.2972 | -87.8288
8.2976 | -87.8283
8.2979 | -87.8278
8.2982 | -87.8277
8.2986 | -87.828
8.2989 | -87.8289
8.2992 | -87.8285
8.2995 | -87.8289
8.2999 | -87.8281
8.3002 | -87.828
8.3005 | -87.8279
8.3009 | -87.827
8.3012 | -87.8262
8.3015 | -87.8258
8.3018 | -87.8254
8.3022 | -87.8254
8.3025 | -87.8253
8.3028 | -87.8252
8.3032 | -87.8248
8.3035 | -87.8244
8.3038 | -87.8238
8.3041 | -87.8225
8.3045 | -87.8221
8.3048 | -87.8221
8.3051 | -87.8217
8.3055 | -87.8216
8.3058 | -87.8212
8.3061 | -87.8209
8.3065 | -87.821
8.3068 | -87.8209
8.3071 | -87.8207
8.3074 | -87.8205
8.3078 | -87.8199
8.3081 | -87.82
8.3084 | -87.8199
8.3088 | -87.8198
8.3091 | -87.8198
8.3094 | -87.8199
8.3097 | -87.8197
8.3101 | -87.8196
8.3104 | -87.8194
8.3107 | -87.819
8.3111 | -87.8187
8.3114 | -87.8182
8.3117 | -87.818
8.312 | -87.8179
8.3124 | -87.8174
8.3127 | -87.817
8.313 | -87.8166
8.3134 | -87.8163
8.3137 | -87.8159
8.314 | -87.8156
8.3144 | -87.8152
8.3147 | -87.815
8.315 | -87.8148
8.3153 | -87.814
8.3157 | -87.813
8.316 | -87.8131
8.3163 | -87.8138
8.3167 | -87.814
8.317 | -87.8144
8.3173 | -87.8138
8.3176 | -87.8134
8.318 | -87.8128
8.3183 | -87.8124
8.3186 | -87.8121
8.319 | -87.8119
8.3193 | -87.8115
8.3196 | -87.8114
8.3199 | -87.8109
8.3203 | -87.811
8.3206 | -87.8111
8.3209 | -87.8104
8.3213 | -87.8098
8.3216 | -87.8096
8.3219 | -87.8098
8.3223 | -87.8095
8.3226 | -87.8094
8.3229 | -87.8093
8.3232 | -87.809
8.3236 | -87.8086
8.3239 | -87.8088
8.3242 | -87.8085
8.3246 | -87.8083
8.3249 | -87.8079
8.3252 | -87.8075
8.3255 | -87.8073
8.3259 | -87.8072
8.3262 | -87.8066
8.3265 | -87.8062
8.3269 | -87.806
8.3272 | -87.8058
8.3275 | -87.8055
8.3278 | -87.8051
8.3282 | -87.8052
8.3285 | -87.8044
8.3288 | -87.8039
8.3292 | -87.8034
8.3295 | -87.8032
8.3298 | -87.8028
8.3302 | -87.8024
8.3305 | -87.8022
8.3308 | -87.8019
8.3311 | -87.8014
8.3315 | -87.8009
8.3318 | -87.801
8.3321 | -87.8007
8.3325 | -87.8007
8.3328 | -87.8003
8.3331 | -87.8001
8.3334 | -87.8002
8.3338 | -87.8001
8.3341 | -87.7998
8.3344 | -87.7997
8.3348 | -87.7994
8.3351 | -87.7988
8.3354 | -87.7984
8.3357 | -87.7986
8.3361 | -87.7986
8.3364 | -87.7986
8.3367 | -87.7978
8.3371 | -87.7969
8.3374 | -87.7963
8.3377 | -87.7954
8.3381 | -87.7951
8.3384 | -87.795
8.3387 | -87.7949
8.339 | -87.7943
8.3394 | -87.7958
8.3397 | -87.7962
8.34 | -87.7959
8.3404 | -87.7956
8.3407 | -87.7953
8.341 | -87.7946
8.3413 | -87.7934
8.3417 | -87.7931
8.342 | -87.7932
8.3423 | -87.7932
8.3427 | -87.7929
8.343 | -87.7926
8.3433 | -87.7925
8.3436 | -87.7924
8.344 | -87.7922
8.3443 | -87.7921
8.3446 | -87.7917
8.345 | -87.7925
8.3453 | -87.7932
8.3456 | -87.7933
8.346 | -87.7935
8.3463 | -87.7928
8.3466 | -87.792
8.3469 | -87.7916
8.3473 | -87.7915
8.3476 | -87.7907
8.3479 | -87.7888
8.3483 | -87.7888
8.3486 | -87.789
8.3489 | -87.7888
8.3492 | -87.7887
8.3496 | -87.7885
8.3499 | -87.7883
8.3502 | -87.7883
8.3506 | -87.7887
8.3509 | -87.7889
8.3512 | -87.7888
8.3515 | -87.7887
8.3519 | -87.788
8.3522 | -87.7874
8.3525 | -87.7871
8.3529 | -87.7869
8.3532 | -87.7863
8.3535 | -87.7855
8.3538 | -87.7853
8.3542 | -87.7851
8.3545 | -87.7846
8.3548 | -87.7839
8.3552 | -87.7834
8.3555 | -87.7829
8.3558 | -87.7824
8.3562 | -87.7811
8.3565 | -87.78
8.3568 | -87.78
8.3571 | -87.7798
8.3575 | -87.7797
8.3578 | -87.779
8.3581 | -87.7784
8.3585 | -87.7781
8.3588 | -87.7781
8.3591 | -87.778
8.3594 | -87.7774
8.3598 | -87.7771
8.3601 | -87.7767
8.3604 | -87.776
8.3608 | -87.7759
8.3611 | -87.7757
8.3614 | -87.7755
8.3617 | -87.7754
8.3621 | -87.7752
8.3624 | -87.775
8.3627 | -87.7747
8.3631 | -87.7749
8.3634 | -87.7752
8.3637 | -87.7746
8.3641 | -87.7744
8.3644 | -87.7745
8.3647 | -87.7743
8.365 | -87.7738
8.3654 | -87.7737
8.3657 | -87.7735
8.366 | -87.7733
8.3664 | -87.7737
8.3667 | -87.7734
8.367 | -87.7729
8.3673 | -87.7727
8.3677 | -87.7731
8.368 | -87.7728
8.3683 | -87.7731
8.3687 | -87.773
8.369 | -87.773
8.3693 | -87.7724
8.3696 | -87.772
8.37 | -87.7719
8.3703 | -87.7718
8.3706 | -87.7719
8.371 | -87.7718
8.3713 | -87.7716
8.3716 | -87.7713
8.372 | -87.7716
8.3723 | -87.7723
8.3726 | -87.7721
8.3729 | -87.7719
8.3733 | -87.7717
8.3736 | -87.7714
8.3739 | -87.771
8.3743 | -87.7708
8.3746 | -87.7707
8.3749 | -87.7704
8.3752 | -87.7697
8.3756 | -87.7697
8.3759 | -87.7694
8.3762 | -87.7692
8.3766 | -87.769
8.3769 | -87.7682
8.3772 | -87.7677
8.3775 | -87.7675
8.3779 | -87.7672
8.3782 | -87.767
8.3785 | -87.7671
8.3789 | -87.7676
8.3792 | -87.7679
8.3795 | -87.7679
8.3799 | -87.7679
8.3802 | -87.7677
8.3805 | -87.7676
8.3808 | -87.7673
8.3812 | -87.7673
8.3815 | -87.7671
8.3818 | -87.7663
8.3822 | -87.766
8.3825 | -87.7663
8.3828 | -87.7667
8.3831 | -87.7668
8.3835 | -87.7662
8.3838 | -87.7654
8.3841 | -87.7651
8.3845 | -87.765
8.3848 | -87.7651
8.3851 | -87.765
8.3854 | -87.7658
8.3858 | -87.7668
8.3861 | -87.7675
8.3864 | -87.7676
8.3868 | -87.7676
8.3871 | -87.7677
8.3874 | -87.7679
8.3878 | -87.768
8.3881 | -87.7684
8.3884 | -87.7684
8.3887 | -87.7684
8.3891 | -87.7681
8.3894 | -87.7672
8.3897 | -87.7668
8.3901 | -87.7663
8.3904 | -87.7658
8.3907 | -87.7656
8.391 | -87.7653
8.3914 | -87.7647
8.3917 | -87.764
8.392 | -87.7641
8.3924 | -87.7641
8.3927 | -87.7641
8.393 | -87.7644
8.3933 | -87.7642
8.3937 | -87.7641
8.394 | -87.7639
8.3943 | -87.7641
8.3947 | -87.7649
8.395 | -87.7649
8.3953 | -87.7649
8.3956 | -87.7646
8.396 | -87.7646
8.3963 | -87.7648
8.3966 | -87.7649
8.397 | -87.7649
8.3973 | -87.7652
8.3976 | -87.7656
8.398 | -87.7662
8.3983 | -87.7667
8.3986 | -87.7664
8.3989 | -87.7667
8.3993 | -87.7666
8.3996 | -87.7663
8.3999 | -87.7661
8.4003 | -87.766
8.4006 | -87.766
8.4009 | -87.7664
8.4012 | -87.7668
8.4016 | -87.7667
8.4019 | -87.7664
8.4022 | -87.7662
8.4026 | -87.7661
8.4029 | -87.766
8.4032 | -87.7658
8.4035 | -87.7654
8.4039 | -87.7656
8.4042 | -87.7664
8.4045 | -87.7662
8.4049 | -87.7659
8.4052 | -87.766
8.4055 | -87.7659
8.4059 | -87.7664
8.4062 | -87.7674
8.4065 | -87.768
8.4068 | -87.7683
8.4072 | -87.7685
8.4075 | -87.7685
8.4078 | -87.7686
8.4082 | -87.7683
8.4085 | -87.7676
8.4088 | -87.7673
8.4091 | -87.7675
8.4095 | -87.7677
8.4098 | -87.7679
8.4101 | -87.7679
8.4105 | -87.7679
8.4108 | -87.7678
8.4111 | -87.768
8.4114 | -87.7675
8.4118 | -87.7675
8.4121 | -87.7678
8.4124 | -87.7679
8.4128 | -87.7689
8.4131 | -87.7696
8.4134 | -87.7697
8.4138 | -87.7705
8.4141 | -87.7703
8.4144 | -87.7703
8.4147 | -87.7706
8.4151 | -87.7706
8.4154 | -87.7704
8.4157 | -87.77
8.4161 | -87.7695
8.4164 | -87.7692
8.4167 | -87.7694
8.417 | -87.7697
8.4174 | -87.77
8.4177 | -87.7696
8.418 | -87.7693
8.4184 | -87.7693
8.4187 | -87.7691
8.419 | -87.769
8.4193 | -87.7691
8.4197 | -87.7692
8.42 | -87.7693
8.4203 | -87.7697
8.4207 | -87.7705
8.421 | -87.7712
8.4213 | -87.7713
8.4217 | -87.7713
8.422 | -87.7714
8.4223 | -87.7714
8.4226 | -87.7725
8.423 | -87.7733
8.4233 | -87.774
8.4236 | -87.774
8.424 | -87.7749
8.4243 | -87.7757
8.4246 | -87.7765
8.4249 | -87.7766
8.4253 | -87.7766
8.4256 | -87.7769
8.4259 | -87.7767
8.4263 | -87.777
8.4266 | -87.7773
8.4269 | -87.7775
8.4272 | -87.7775
8.4276 | -87.7777
8.4279 | -87.7779
8.4282 | -87.7784
8.4286 | -87.7791
8.4289 | -87.7794
8.4292 | -87.7797
8.4296 | -87.78
8.4299 | -87.7805
8.4302 | -87.7805
8.4305 | -87.7808
8.4309 | -87.7811
8.4312 | -87.7818
8.4315 | -87.7821
8.4319 | -87.7823
8.4322 | -87.7824
8.4325 | -87.7834
8.4328 | -87.7845
8.4332 | -87.7854
8.4335 | -87.7856
8.4338 | -87.7856
8.4342 | -87.7851
8.4345 | -87.7856
8.4348 | -87.7871
8.4351 | -87.7877
8.4355 | -87.7875
8.4358 | -87.7877
8.4361 | -87.7879
8.4365 | -87.7883
8.4368 | -87.7888
8.4371 | -87.7898
8.4375 | -87.7901
8.4378 | -87.7901
8.4381 | -87.7902
8.4384 | -87.7906
8.4388 | -87.7914
8.4391 | -87.7921
8.4394 | -87.7924
8.4398 | -87.7924
8.4401 | -87.793
8.4404 | -87.7932
8.4407 | -87.7932
8.4411 | -87.793
8.4414 | -87.7927
8.4417 | -87.7929
8.4421 | -87.7939
8.4424 | -87.7944
8.4427 | -87.7945
8.443 | -87.7946
8.4434 | -87.7947
8.4437 | -87.7951
8.444 | -87.796
8.4444 | -87.7965
8.4447 | -87.7967
8.445 | -87.7975
8.4453 | -87.7986
8.4457 | -87.7995
8.446 | -87.8003
8.4463 | -87.8009
8.4467 | -87.8018
8.447 | -87.8028
8.4473 | -87.8031
8.4477 | -87.8036
8.448 | -87.8038
8.4483 | -87.8044
8.4486 | -87.8047
8.449 | -87.8054
8.4493 | -87.8063
8.4496 | -87.807
8.45 | -87.8076
8.4503 | -87.8081
8.4506 | -87.8087
8.4509 | -87.8092
8.4513 | -87.8095
8.4516 | -87.8101
8.4519 | -87.8111
8.4523 | -87.8117
8.4526 | -87.8121
8.4529 | -87.8129
8.4532 | -87.8129
8.4536 | -87.8128
8.4539 | -87.8134
8.4542 | -87.814
8.4546 | -87.8149
8.4549 | -87.8152
8.4552 | -87.8158
8.4556 | -87.8165
8.4559 | -87.8171
8.4562 | -87.8183
8.4565 | -87.8187
8.4569 | -87.8192
8.4572 | -87.8208
8.4575 | -87.822
8.4579 | -87.8226
8.4582 | -87.8233
8.4585 | -87.823
8.4588 | -87.8232
8.4592 | -87.8236
8.4595 | -87.8236
8.4598 | -87.824
8.4602 | -87.824
8.4605 | -87.8245
8.4608 | -87.8261
8.4611 | -87.8268
8.4615 | -87.8278
8.4618 | -87.8274
8.4621 | -87.8277
8.4625 | -87.8285
8.4628 | -87.8293
8.4631 | -87.8291
8.4635 | -87.829
8.4638 | -87.8296
8.4641 | -87.83
8.4644 | -87.8303
8.4648 | -87.8317
8.4651 | -87.8326
8.4654 | -87.833
8.4658 | -87.8333
8.4661 | -87.8335
8.4664 | -87.8342
8.4667 | -87.8346
8.4671 | -87.8349
8.4674 | -87.8356
8.4677 | -87.8359
8.4681 | -87.8359
8.4684 | -87.8364
8.4687 | -87.8367
8.469 | -87.837
8.4694 | -87.8378
8.4697 | -87.8383
8.47 | -87.8386
8.4704 | -87.8387
8.4707 | -87.8394
8.471 | -87.8396
8.4714 | -87.8399
8.4717 | -87.8403
8.472 | -87.84
8.4723 | -87.8396
8.4727 | -87.8401
8.473 | -87.8404
8.4733 | -87.8407
8.4737 | -87.8407
8.474 | -87.8403
8.4743 | -87.8408
8.4746 | -87.8416
8.475 | -87.8418
8.4753 | -87.8422
8.4756 | -87.8426
8.476 | -87.8432
8.4763 | -87.8445
8.4766 | -87.8457
8.4769 | -87.8464
8.4773 | -87.8461
8.4776 | -87.8461
8.4779 | -87.8467
8.4783 | -87.8471
8.4786 | -87.8474
8.4789 | -87.8471
8.4793 | -87.8473
8.4796 | -87.8478
8.4799 | -87.8481
8.4802 | -87.8486
8.4806 | -87.8493
8.4809 | -87.8501
8.4812 | -87.8502
8.4816 | -87.8512
8.4819 | -87.852
8.4822 | -87.8529
8.4825 | -87.8534
8.4829 | -87.8544
8.4832 | -87.8553
8.4835 | -87.8557
8.4839 | -87.8562
8.4842 | -87.857
8.4845 | -87.8575
8.4848 | -87.8578
8.4852 | -87.8588
8.4855 | -87.8593
8.4858 | -87.8593
8.4862 | -87.8601
8.4865 | -87.8608
8.4868 | -87.8614
8.4871 | -87.8624
8.4875 | -87.863
8.4878 | -87.8635
8.4881 | -87.8639
8.4885 | -87.8641
8.4888 | -87.8647
8.4891 | -87.8652
8.4895 | -87.8656
8.4898 | -87.8668
8.4901 | -87.8674
8.4904 | -87.8679
8.4908 | -87.8687
8.4911 | -87.8694
8.4914 | -87.8701
8.4918 | -87.871
8.4921 | -87.872
8.4924 | -87.8728
8.4927 | -87.8739
8.4931 | -87.8758
8.4934 | -87.8769
8.4937 | -87.8773
8.4941 | -87.878
8.4944 | -87.8786
8.4947 | -87.8791
8.4951 | -87.8802
8.4954 | -87.8807
8.4957 | -87.8814
8.496 | -87.882
8.4964 | -87.8829
8.4967 | -87.8843
8.497 | -87.8859
8.4974 | -87.8875
8.4977 | -87.8884
8.498 | -87.8891
8.4983 | -87.8897
8.4987 | -87.8905
8.499 | -87.892
8.4993 | -87.8933
8.4997 | -87.8952
8.5 | -87.8971
8.5003 | -87.8987
8.5006 | -87.8994
8.501 | -87.9001
8.5013 | -87.9008
8.5016 | -87.9018
8.502 | -87.9026
8.5023 | -87.9031
8.5026 | -87.9038
8.5029 | -87.9045
8.5033 | -87.905
8.5036 | -87.9057
8.5039 | -87.9062
8.5043 | -87.907
8.5046 | -87.9077
8.5049 | -87.9081
8.5053 | -87.9081
8.5056 | -87.9086
8.5059 | -87.9093
8.5062 | -87.91
8.5066 | -87.9107
8.5069 | -87.9114
8.5072 | -87.9122
8.5076 | -87.9134
8.5079 | -87.9142
8.5082 | -87.9158
8.5085 | -87.917
8.5089 | -87.9183
8.5092 | -87.919
8.5095 | -87.9197
8.5099 | -87.9206
8.5102 | -87.9214
8.5105 | -87.9225
8.5108 | -87.9229
8.5112 | -87.9236
8.5115 | -87.9248
8.5118 | -87.9258
8.5122 | -87.9265
8.5125 | -87.9274
8.5128 | -87.9277
8.5132 | -87.9283
8.5135 | -87.9292
8.5138 | -87.9302
8.5141 | -87.9315
8.5145 | -87.9326
8.5148 | -87.9334
8.5151 | -87.934
8.5155 | -87.9358
8.5158 | -87.9371
8.5161 | -87.938
8.5164 | -87.9387
8.5168 | -87.9392
8.5171 | -87.9396
8.5174 | -87.94
8.5178 | -87.941
8.5181 | -87.9424
8.5184 | -87.9428
8.5187 | -87.9433
8.5191 | -87.9441
8.5194 | -87.9442
8.5197 | -87.9452
8.5201 | -87.9459
8.5204 | -87.9467
8.5207 | -87.947
8.5211 | -87.9477
8.5214 | -87.9485
8.5217 | -87.9493
8.522 | -87.9504
8.5224 | -87.9519
8.5227 | -87.9531
8.523 | -87.9536
8.5234 | -87.9544
8.5237 | -87.9553
8.524 | -87.9563
8.5243 | -87.957
8.5247 | -87.9577
8.525 | -87.9587
8.5253 | -87.9598
8.5257 | -87.9606
8.526 | -87.9617
8.5263 | -87.9634
8.5266 | -87.9643
8.527 | -87.9646
8.5273 | -87.9654
8.5276 | -87.967
8.528 | -87.9678
8.5283 | -87.9687
8.5286 | -87.9694
8.529 | -87.9698
8.5293 | -87.9697
8.5296 | -87.9695
8.5299 | -87.9704
8.5303 | -87.9717
8.5306 | -87.9726
8.5309 | -87.9732
8.5313 | -87.9742
8.5316 | -87.9751
8.5319 | -87.9757
8.5322 | -87.9765
8.5326 | -87.9771
8.5329 | -87.9778
8.5332 | -87.9786
8.5336 | -87.9794
8.5339 | -87.9804
8.5342 | -87.9814
8.5345 | -87.9822
8.5349 | -87.9831
8.5352 | -87.9837
8.5355 | -87.9844
8.5359 | -87.9854
8.5362 | -87.9864
8.5365 | -87.9873
8.5368 | -87.9881
8.5372 | -87.9894
8.5375 | -87.9905
8.5378 | -87.9915
8.5382 | -87.9922
8.5385 | -87.9932
8.5388 | -87.9942
8.5392 | -87.9948
8.5395 | -87.9957
8.5398 | -87.9966
8.5401 | -87.9975
8.5405 | -87.9982
8.5408 | -87.999
8.5411 | -87.9997
8.5415 | -88.0009
8.5418 | -88.0022
8.5421 | -88.0033
8.5424 | -88.0044
8.5428 | -88.0052
8.5431 | -88.006
8.5434 | -88.007
8.5438 | -88.0079
8.5441 | -88.009
8.5444 | -88.01
8.5448 | -88.0109
8.5451 | -88.0117
8.5454 | -88.0129
8.5457 | -88.014
8.5461 | -88.0146
8.5464 | -88.0155
8.5467 | -88.017
8.5471 | -88.0179
8.5474 | -88.0189
8.5477 | -88.02
8.548 | -88.021
8.5484 | -88.0216
8.5487 | -88.0224
8.549 | -88.0232
8.5494 | -88.0237
8.5497 | -88.024
8.55 | -88.0241
8.5503 | -88.0246
8.5507 | -88.0256
8.551 | -88.0264
8.5513 | -88.027
8.5517 | -88.0278
8.552 | -88.0291
8.5523 | -88.0303
8.5526 | -88.0312
8.553 | -88.0326
8.5533 | -88.0345
8.5536 | -88.0355
8.554 | -88.0368
8.5543 | -88.0376
8.5546 | -88.0388
8.555 | -88.0397
8.5553 | -88.0409
8.5556 | -88.0421
8.5559 | -88.0431
8.5563 | -88.0443
8.5566 | -88.0449
8.5569 | -88.0449
8.5573 | -88.0453
8.5576 | -88.0461
8.5579 | -88.0472
8.5582 | -88.0485
8.5586 | -88.0495
8.5589 | -88.0507
8.5592 | -88.0533
8.5596 | -88.0554
8.5599 | -88.0574
8.5602 | -88.0597
8.5605 | -88.0612
8.5609 | -88.0617
8.5612 | -88.0624
8.5615 | -88.0632
8.5619 | -88.0639
8.5622 | -88.065
8.5625 | -88.066
8.5629 | -88.0677
8.5632 | -88.0685
8.5635 | -88.0693
8.5638 | -88.0702
8.5642 | -88.0712
8.5645 | -88.0719
8.5648 | -88.0718
8.5652 | -88.0719
8.5655 | -88.0723
8.5658 | -88.0734
8.5661 | -88.0744
8.5665 | -88.0753
8.5668 | -88.0762
8.5671 | -88.0768
8.5675 | -88.0776
8.5678 | -88.0786
8.5681 | -88.0796
8.5684 | -88.0808
8.5688 | -88.0821
8.5691 | -88.0834
8.5694 | -88.0843
8.5698 | -88.0852
8.5701 | -88.0863
8.5704 | -88.0871
8.5708 | -88.0876
8.5711 | -88.088
8.5714 | -88.0885
8.5717 | -88.0895
8.5721 | -88.0905
8.5724 | -88.0912
8.5727 | -88.0922
8.5731 | -88.0931
8.5734 | -88.0941
8.5737 | -88.0954
8.574 | -88.0965
8.5744 | -88.0975
8.5747 | -88.0986
8.575 | -88.0995
8.5754 | -88.1007
8.5757 | -88.1017
8.576 | -88.103
8.5763 | -88.104
8.5767 | -88.1048
8.577 | -88.1055
8.5773 | -88.1065
8.5777 | -88.1076
8.578 | -88.1088
8.5783 | -88.1106
8.5787 | -88.1119
8.579 | -88.1133
8.5793 | -88.1146
8.5796 | -88.1155
8.58 | -88.1169
8.5803 | -88.1195
8.5806 | -88.1209
8.581 | -88.1221
8.5813 | -88.1228
8.5816 | -88.1237
8.5819 | -88.1249
8.5823 | -88.1261
8.5826 | -88.1269
8.5829 | -88.128
8.5833 | -88.1297
8.5836 | -88.131
8.5839 | -88.1323
8.5842 | -88.1334
8.5846 | -88.1342
8.5849 | -88.1352
8.5852 | -88.1363
8.5856 | -88.1372
8.5859 | -88.1386
8.5862 | -88.1399
8.5866 | -88.1412
8.5869 | -88.1424
8.5872 | -88.1441
8.5875 | -88.1452
8.5879 | -88.1467
8.5882 | -88.1479
8.5885 | -88.1489
8.5889 | -88.1499
8.5892 | -88.151
8.5895 | -88.152
8.5898 | -88.1532
8.5902 | -88.1541
8.5905 | -88.1553
8.5908 | -88.1568
8.5912 | -88.1576
8.5915 | -88.1581
8.5918 | -88.1591
8.5921 | -88.1603
8.5925 | -88.1612
8.5928 | -88.1622
8.5931 | -88.1629
8.5935 | -88.1638
8.5938 | -88.1654
8.5941 | -88.1668
8.5944 | -88.168
8.5948 | -88.1693
8.5951 | -88.1707
8.5954 | -88.1718
8.5958 | -88.1728
8.5961 | -88.174
8.5964 | -88.1757
8.5968 | -88.1771
8.5971 | -88.1779
8.5974 | -88.1785
8.5977 | -88.1795
8.5981 | -88.1806
8.5984 | -88.1813
8.5987 | -88.1817
8.5991 | -88.1827
8.5994 | -88.1838
8.5997 | -88.1849
8.6 | -88.1858
8.6004 | -88.1868
8.6007 | -88.1879
8.601 | -88.189
8.6014 | -88.1904
8.6017 | -88.1925
8.602 | -88.1943
8.6023 | -88.1958
8.6027 | -88.1968
8.603 | -88.1979
8.6033 | -88.1999
8.6037 | -88.2005
8.604 | -88.2027
8.6043 | -88.2045
8.6047 | -88.2055
8.605 | -88.2064
8.6053 | -88.2072
8.6056 | -88.2081
8.606 | -88.2091
8.6063 | -88.2102
8.6066 | -88.2115
8.607 | -88.2129
8.6073 | -88.2141
8.6076 | -88.2151
8.6079 | -88.2162
8.6083 | -88.2172
8.6086 | -88.2182
8.6089 | -88.2194
8.6093 | -88.2206
8.6096 | -88.2218
8.6099 | -88.224
8.6102 | -88.2257
8.6106 | -88.2266
8.6109 | -88.2277
8.6112 | -88.2289
8.6116 | -88.2298
8.6119 | -88.2307
8.6122 | -88.2318
8.6126 | -88.2329
8.6129 | -88.2341
8.6132 | -88.2339
8.6135 | -88.2349
8.6139 | -88.236
8.6142 | -88.2369
8.6145 | -88.238
8.6149 | -88.2395
8.6152 | -88.2415
8.6155 | -88.2433
8.6158 | -88.2446
8.6162 | -88.246
8.6165 | -88.2473
8.6168 | -88.2483
8.6172 | -88.2494
8.6175 | -88.2502
8.6178 | -88.2513
8.6181 | -88.2525
8.6185 | -88.2534
8.6188 | -88.2538
8.6191 | -88.2543
8.6195 | -88.2553
8.6198 | -88.2565
8.6201 | -88.2575
8.6205 | -88.2588
8.6208 | -88.2601
8.6211 | -88.2619
8.6214 | -88.263
8.6218 | -88.2638
8.6221 | -88.2646
8.6224 | -88.2655
8.6228 | -88.2666
8.6231 | -88.2678
8.6234 | -88.269
8.6237 | -88.2703
8.6241 | -88.2716
8.6244 | -88.2726
8.6247 | -88.2735
8.6251 | -88.2745
8.6254 | -88.2755
8.6257 | -88.2769
8.626 | -88.2778
8.6264 | -88.2788
8.6267 | -88.28
8.627 | -88.2818
8.6274 | -88.2835
8.6277 | -88.2847
8.628 | -88.2859
8.6284 | -88.2872
8.6287 | -88.2888
8.629 | -88.2901
8.6293 | -88.2913
8.6297 | -88.2923
8.63 | -88.2932
8.6303 | -88.294
8.6307 | -88.295
8.631 | -88.2965
8.6313 | -88.2975
8.6316 | -88.2985
8.632 | -88.2999
8.6323 | -88.3011
8.6326 | -88.302
8.633 | -88.3028
8.6333 | -88.3043
8.6336 | -88.3058
8.6339 | -88.3079
8.6343 | -88.31
8.6346 | -88.3111
8.6349 | -88.3122
8.6353 | -88.3134
8.6356 | -88.3151
8.6359 | -88.3165
8.6363 | -88.317
8.6366 | -88.3181
8.6369 | -88.3194
8.6372 | -88.3206
8.6376 | -88.3217
8.6379 | -88.3227
8.6382 | -88.3238
8.6386 | -88.326
8.6389 | -88.328
8.6392 | -88.3294
8.6395 | -88.3307
8.6399 | -88.3318
8.6402 | -88.3328
8.6405 | -88.3339
8.6409 | -88.3351
8.6412 | -88.3362
8.6415 | -88.3375
8.6418 | -88.3388
8.6422 | -88.3399
8.6425 | -88.3409
8.6428 | -88.342
8.6432 | -88.3435
8.6435 | -88.3451
8.6438 | -88.3459
8.6441 | -88.3469
8.6445 | -88.3483
8.6448 | -88.3494
8.6451 | -88.351
8.6455 | -88.3526
8.6458 | -88.3544
8.6461 | -88.3561
8.6465 | -88.3579
8.6468 | -88.3589
8.6471 | -88.3602
8.6474 | -88.3613
8.6478 | -88.3624
8.6481 | -88.3635
8.6484 | -88.3647
8.6488 | -88.3658
8.6491 | -88.367
8.6494 | -88.3689
8.6497 | -88.371
8.6501 | -88.3726
8.6504 | -88.3742
8.6507 | -88.3762
8.6511 | -88.3775
8.6514 | -88.3793
8.6517 | -88.3807
8.652 | -88.3819
8.6524 | -88.3835
8.6527 | -88.3852
8.653 | -88.3865
8.6534 | -88.3876
8.6537 | -88.3883
8.654 | -88.3895
8.6544 | -88.3913
8.6547 | -88.3924
8.655 | -88.3936
8.6553 | -88.3951
8.6557 | -88.3968
8.656 | -88.398
8.6563 | -88.3993
8.6567 | -88.4003
8.657 | -88.4011
8.6573 | -88.4018
8.6576 | -88.4035
8.658 | -88.405
8.6583 | -88.4069
8.6586 | -88.4081
8.659 | -88.4097
8.6593 | -88.4107
8.6596 | -88.4119
8.6599 | -88.4126
8.6603 | -88.4134
8.6606 | -88.4145
8.6609 | -88.4152
8.6613 | -88.4161
8.6616 | -88.4174
8.6619 | -88.4192
8.6623 | -88.4212
8.6626 | -88.4225
8.6629 | -88.424
8.6632 | -88.4256
8.6636 | -88.4267
8.6639 | -88.4273
8.6642 | -88.4285
8.6646 | -88.43
8.6649 | -88.4312
8.6652 | -88.4323
8.6655 | -88.4334
8.6659 | -88.4346
8.6662 | -88.4358
8.6665 | -88.4367
8.6669 | -88.4378
8.6672 | -88.4389
8.6675 | -88.4399
8.6678 | -88.4409
8.6682 | -88.4406
8.6685 | -88.4411
8.6688 | -88.4429
8.6692 | -88.4446
8.6695 | -88.4459
8.6698 | -88.447
8.6702 | -88.4486
8.6705 | -88.4497
8.6708 | -88.4508
8.6711 | -88.4518
8.6715 | -88.4527
8.6718 | -88.4529
8.6721 | -88.4537
8.6725 | -88.4551
8.6728 | -88.4567
8.6731 | -88.458
8.6734 | -88.46
8.6738 | -88.4609
8.6741 | -88.4619
8.6744 | -88.4636
8.6748 | -88.4653
8.6751 | -88.4666
8.6754 | -88.4673
8.6757 | -88.4688
8.6761 | -88.4702
8.6764 | -88.471
8.6767 | -88.4721
8.6771 | -88.4737
8.6774 | -88.4757
8.6777 | -88.4773
8.6781 | -88.4782
8.6784 | -88.4788
8.6787 | -88.4798
8.679 | -88.4807
8.6794 | -88.4818
8.6797 | -88.4833
8.68 | -88.484
8.6804 | -88.4846
8.6807 | -88.4854
8.681 | -88.4865
8.6813 | -88.4877
8.6817 | -88.4886
8.682 | -88.4894
8.6823 | -88.4903
8.6827 | -88.4913
8.683 | -88.4925
8.6833 | -88.4936
8.6836 | -88.4947
8.684 | -88.4962
8.6843 | -88.4973
8.6846 | -88.4985
8.685 | -88.4998
8.6853 | -88.5009
8.6856 | -88.5021
8.686 | -88.5031
8.6863 | -88.5041
8.6866 | -88.5049
8.6869 | -88.5059
8.6873 | -88.5077
8.6876 | -88.509
8.6879 | -88.5102
8.6883 | -88.5113
8.6886 | -88.5124
8.6889 | -88.5138
8.6892 | -88.5148
8.6896 | -88.5156
8.6899 | -88.5168
8.6902 | -88.5182
8.6906 | -88.5194
8.6909 | -88.5206
8.6912 | -88.5219
8.6915 | -88.5233
8.6919 | -88.5245
8.6922 | -88.5254
8.6925 | -88.5261
8.6929 | -88.5273
8.6932 | -88.5285
8.6935 | -88.5293
8.6938 | -88.5299
8.6942 | -88.5315
8.6945 | -88.5339
8.6948 | -88.5355
8.6952 | -88.5363
8.6955 | -88.5372
8.6958 | -88.5384
8.6962 | -88.539
8.6965 | -88.5397
8.6968 | -88.541
8.6971 | -88.5424
8.6975 | -88.544
8.6978 | -88.5452
8.6981 | -88.5464
8.6985 | -88.5476
8.6988 | -88.549
8.6991 | -88.5502
8.6994 | -88.5514
8.6998 | -88.5525
8.7001 | -88.5536
8.7004 | -88.5547
8.7008 | -88.5561
8.7011 | -88.5571
8.7014 | -88.5583
8.7017 | -88.5593
8.7021 | -88.5601
8.7024 | -88.5613
8.7027 | -88.5622
8.7031 | -88.5631
8.7034 | -88.5647
8.7037 | -88.5661
8.7041 | -88.5668
8.7044 | -88.568
8.7047 | -88.5693
8.705 | -88.5709
8.7054 | -88.5722
8.7057 | -88.5727
8.706 | -88.5735
8.7064 | -88.5746
8.7067 | -88.5761
8.707 | -88.5775
8.7073 | -88.5789
8.7077 | -88.5803
8.708 | -88.5816
8.7083 | -88.5823
8.7087 | -88.5834
8.709 | -88.5849
8.7093 | -88.5862
8.7096 | -88.5872
8.71 | -88.5894
8.7103 | -88.5918
8.7106 | -88.5927
8.711 | -88.5935
8.7113 | -88.5946
8.7116 | -88.5959
8.712 | -88.5966
8.7123 | -88.597
8.7126 | -88.5988
8.7129 | -88.6005
8.7133 | -88.6018
8.7136 | -88.6034
8.7139 | -88.6059
8.7143 | -88.6074
8.7146 | -88.6084
8.7149 | -88.6096
8.7152 | -88.6108
8.7156 | -88.6129
8.7159 | -88.6143
8.7162 | -88.6162
8.7166 | -88.6184
8.7169 | -88.6199
8.7172 | -88.6209
8.7175 | -88.6224
8.7179 | -88.6237
8.7182 | -88.6258
8.7185 | -88.6284
8.7189 | -88.6307
8.7192 | -88.6326
8.7195 | -88.6337
8.7199 | -88.6346
8.7202 | -88.6369
8.7205 | -88.639
8.7208 | -88.6407
8.7212 | -88.6427
8.7215 | -88.6443
8.7218 | -88.6452
8.7222 | -88.6468
8.7225 | -88.6479
8.7228 | -88.6486
8.7231 | -88.6504
8.7235 | -88.6521
8.7238 | -88.653
8.7241 | -88.654
8.7245 | -88.6548
8.7248 | -88.6553
8.7251 | -88.6563
8.7254 | -88.6574
8.7258 | -88.6584
8.7261 | -88.6596
8.7264 | -88.6607
8.7268 | -88.662
8.7271 | -88.6632
8.7274 | -88.6642
8.7278 | -88.6656
8.7281 | -88.6673
8.7284 | -88.6687
8.7287 | -88.6691
8.7291 | -88.67
8.7294 | -88.6715
8.7297 | -88.6729
8.7301 | -88.6737
8.7304 | -88.6748
8.7307 | -88.6761
8.731 | -88.6776
8.7314 | -88.6785
8.7317 | -88.6794
8.732 | -88.681
8.7324 | -88.6828
8.7327 | -88.6838
8.733 | -88.6841
8.7333 | -88.6844
8.7337 | -88.6852
8.734 | -88.6863
8.7343 | -88.6874
8.7347 | -88.6887
8.735 | -88.69
8.7353 | -88.6908
8.7356 | -88.6915
8.736 | -88.6927
8.7363 | -88.6939
8.7366 | -88.6953
8.737 | -88.6964
8.7373 | -88.6975
8.7376 | -88.6981
8.738 | -88.6989
8.7383 | -88.6997
8.7386 | -88.7007
8.7389 | -88.7023
8.7393 | -88.7043
8.7396 | -88.7065
8.7399 | -88.7075
8.7403 | -88.7079
8.7406 | -88.709
8.7409 | -88.7103
8.7412 | -88.7114
8.7416 | -88.7125
8.7419 | -88.7129
8.7422 | -88.7134
8.7426 | -88.7156
8.7429 | -88.7168
8.7432 | -88.7174
8.7435 | -88.7186
8.7439 | -88.7201
8.7442 | -88.7216
8.7445 | -88.7232
8.7449 | -88.7245
8.7452 | -88.7255
8.7455 | -88.7267
8.7459 | -88.7278
8.7462 | -88.7291
8.7465 | -88.7301
8.7468 | -88.7316
8.7472 | -88.7328
8.7475 | -88.7335
8.7478 | -88.7347
8.7482 | -88.7358
8.7485 | -88.7364
8.7488 | -88.7371
8.7491 | -88.7381
8.7495 | -88.7398
8.7498 | -88.7421
8.7501 | -88.7432
8.7505 | -88.7446
8.7508 | -88.7459
8.7511 | -88.7472
8.7514 | -88.7485
8.7518 | -88.7508
8.7521 | -88.7518
8.7524 | -88.7528
8.7528 | -88.7539
8.7531 | -88.7553
8.7534 | -88.7574
8.7538 | -88.7587
8.7541 | -88.7594
8.7544 | -88.7602
8.7547 | -88.7614
8.7551 | -88.7625
8.7554 | -88.7632
8.7557 | -88.7637
8.7561 | -88.7644
8.7564 | -88.7653
8.7567 | -88.766
8.757 | -88.7666
8.7574 | -88.7677
8.7577 | -88.7688
8.758 | -88.7699
8.7584 | -88.7709
8.7587 | -88.7717
8.759 | -88.7728
8.7593 | -88.7738
8.7597 | -88.7747
8.76 | -88.7759
8.7603 | -88.7771
8.7607 | -88.7776
8.761 | -88.7784
8.7613 | -88.7794
8.7617 | -88.7805
8.762 | -88.7815
8.7623 | -88.7824
8.7626 | -88.7833
8.763 | -88.7848
8.7633 | -88.7862
8.7636 | -88.7879
8.764 | -88.7904
8.7643 | -88.7917
8.7646 | -88.7925
8.7649 | -88.7936
8.7653 | -88.7946
8.7656 | -88.7961
8.7659 | -88.7972
8.7663 | -88.7977
8.7666 | -88.7983
8.7669 | -88.7991
8.7672 | -88.8002
8.7676 | -88.8014
8.7679 | -88.8023
8.7682 | -88.8033
8.7686 | -88.8042
8.7689 | -88.8053
8.7692 | -88.8072
8.7696 | -88.8089
8.7699 | -88.8101
8.7702 | -88.8122
8.7705 | -88.8134
8.7709 | -88.8149
8.7712 | -88.8162
8.7715 | -88.8168
8.7719 | -88.8178
8.7722 | -88.8189
8.7725 | -88.82
8.7728 | -88.8206
8.7732 | -88.8204
8.7735 | -88.8208
8.7738 | -88.8221
8.7742 | -88.8229
8.7745 | -88.8236
8.7748 | -88.8245
8.7751 | -88.8256
8.7755 | -88.8267
8.7758 | -88.8283
8.7761 | -88.8293
8.7765 | -88.8315
8.7768 | -88.8325
8.7771 | -88.8339
8.7774 | -88.8352
8.7778 | -88.8361
8.7781 | -88.8373
8.7784 | -88.838
8.7788 | -88.8389
8.7791 | -88.8399
8.7794 | -88.8409
8.7798 | -88.8418
8.7801 | -88.843
8.7804 | -88.8443
8.7807 | -88.8459
8.7811 | -88.8472
8.7814 | -88.8483
8.7817 | -88.8493
8.7821 | -88.85
8.7824 | -88.8507
8.7827 | -88.8519
8.783 | -88.8526
8.7834 | -88.8531
8.7837 | -88.8543
8.784 | -88.8554
8.7844 | -88.8577
8.7847 | -88.8592
8.785 | -88.86
8.7853 | -88.8612
8.7857 | -88.8643
8.786 | -88.867
8.7863 | -88.8682
8.7867 | -88.8689
8.787 | -88.8695
8.7873 | -88.8704
8.7877 | -88.8704
8.788 | -88.8716
8.7883 | -88.8724
8.7886 | -88.8733
8.789 | -88.8747
8.7893 | -88.8752
8.7896 | -88.8767
8.79 | -88.8783
8.7903 | -88.8793
8.7906 | -88.8799
8.7909 | -88.881
8.7913 | -88.8821
8.7916 | -88.8831
8.7919 | -88.8841
8.7923 | -88.8852
8.7926 | -88.8868
8.7929 | -88.8878
8.7932 | -88.8889
8.7936 | -88.8902
8.7939 | -88.8913
8.7942 | -88.8923
8.7946 | -88.8932
8.7949 | -88.8946
8.7952 | -88.8956
8.7956 | -88.8966
8.7959 | -88.8974
8.7962 | -88.8982
8.7965 | -88.899
8.7969 | -88.8998
8.7972 | -88.9009
8.7975 | -88.9019
8.7979 | -88.9028
8.7982 | -88.9037
8.7985 | -88.9047
8.7988 | -88.9058
8.7992 | -88.9069
8.7995 | -88.9077
8.7998 | -88.9087
8.8002 | -88.9094
8.8005 | -88.91
8.8008 | -88.91
8.8011 | -88.9101
8.8015 | -88.9107
8.8018 | -88.9126
8.8021 | -88.9148
8.8025 | -88.9152
8.8028 | -88.915
8.8031 | -88.9158
8.8035 | -88.9164
8.8038 | -88.9169
8.8041 | -88.9182
8.8044 | -88.9192
8.8048 | -88.9201
8.8051 | -88.9205
8.8054 | -88.9214
8.8058 | -88.9224
8.8061 | -88.9231
8.8064 | -88.9239
8.8067 | -88.9248
8.8071 | -88.9257
8.8074 | -88.9267
8.8077 | -88.9282
8.8081 | -88.9296
8.8084 | -88.9307
8.8087 | -88.9323
8.809 | -88.9334
8.8094 | -88.9341
8.8097 | -88.9351
8.81 | -88.9366
8.8104 | -88.9384
8.8107 | -88.9392
8.811 | -88.9403
8.8114 | -88.9412
8.8117 | -88.9422
8.812 | -88.943
8.8123 | -88.9445
8.8127 | -88.9455
8.813 | -88.9467
8.8133 | -88.9475
8.8137 | -88.9483
8.814 | -88.9497
8.8143 | -88.9502
8.8146 | -88.9507
8.815 | -88.9516
8.8153 | -88.9525
8.8156 | -88.9535
8.816 | -88.9543
8.8163 | -88.9548
8.8166 | -88.9553
8.8169 | -88.9562
8.8173 | -88.9574
8.8176 | -88.9587
8.8179 | -88.9607
8.8183 | -88.9619
8.8186 | -88.9627
8.8189 | -88.9639
8.8193 | -88.9649
8.8196 | -88.9657
8.8199 | -88.9673
8.8202 | -88.9687
8.8206 | -88.9699
8.8209 | -88.9709
8.8212 | -88.9717
8.8216 | -88.9734
8.8219 | -88.9747
8.8222 | -88.9762
8.8225 | -88.9775
8.8229 | -88.9784
8.8232 | -88.9793
8.8235 | -88.9796
8.8239 | -88.9804
8.8242 | -88.9823
8.8245 | -88.9842
8.8248 | -88.9851
8.8252 | -88.9863
8.8255 | -88.9875
8.8258 | -88.9884
8.8262 | -88.9891
8.8265 | -88.9902
8.8268 | -88.9919
8.8271 | -88.993
8.8275 | -88.9939
8.8278 | -88.995
8.8281 | -88.9962
8.8285 | -88.9973
8.8288 | -88.9973
8.8291 | -88.9959
8.8295 | -88.9961
8.8298 | -88.9972
8.8301 | -88.9975
8.8304 | -88.9976
8.8308 | -88.9983
8.8311 | -88.9991
8.8314 | -88.9996
8.8318 | -89.0003
8.8321 | -89.0013
8.8324 | -89.0023
8.8327 | -89.0034
8.8331 | -89.0043
8.8334 | -89.0051
8.8337 | -89.0056
8.8341 | -89.0065
8.8344 | -89.007
8.8347 | -89.0074
8.835 | -89.0081
8.8354 | -89.0086
8.8357 | -89.0093
8.836 | -89.011
8.8364 | -89.0121
8.8367 | -89.0129
8.837 | -89.0138
8.8374 | -89.0149
8.8377 | -89.0155
8.838 | -89.0163
8.8383 | -89.0171
8.8387 | -89.0179
8.839 | -89.0187
8.8393 | -89.0193
8.8397 | -89.0199
8.84 | -89.0209
8.8403 | -89.022
8.8406 | -89.023
8.841 | -89.0241
8.8413 | -89.0254
8.8416 | -89.0265
8.842 | -89.0276
8.8423 | -89.0287
8.8426 | -89.0299
8.8429 | -89.0309
8.8433 | -89.0316
8.8436 | -89.0322
8.8439 | -89.033
8.8443 | -89.0339
8.8446 | -89.0352
8.8449 | -89.0356
8.8453 | -89.0345
8.8456 | -89.0339
8.8459 | -89.0347
8.8462 | -89.0355
8.8466 | -89.0363
8.8469 | -89.0377
8.8472 | -89.0387
8.8476 | -89.0395
8.8479 | -89.0409
8.8482 | -89.0421
8.8485 | -89.0428
8.8489 | -89.0438
8.8492 | -89.0451
8.8495 | -89.0463
8.8499 | -89.0473
8.8502 | -89.0478
8.8505 | -89.0477
8.8508 | -89.0485
8.8512 | -89.0498
8.8515 | -89.0506
8.8518 | -89.0514
8.8522 | -89.0523
8.8525 | -89.0533
8.8528 | -89.0538
8.8532 | -89.0546
8.8535 | -89.0555
8.8538 | -89.0561
8.8541 | -89.0586
8.8545 | -89.0603
8.8548 | -89.0615
8.8551 | -89.0615
8.8555 | -89.0619
8.8558 | -89.0626
8.8561 | -89.0637
8.8564 | -89.0644
8.8568 | -89.0653
8.8571 | -89.0664
8.8574 | -89.0672
8.8578 | -89.0684
8.8581 | -89.0695
8.8584 | -89.0703
8.8587 | -89.0715
8.8591 | -89.0725
8.8594 | -89.0732
8.8597 | -89.0741
8.8601 | -89.0749
8.8604 | -89.0762
8.8607 | -89.0776
8.8611 | -89.0791
8.8614 | -89.0809
8.8617 | -89.082
8.862 | -89.0829
8.8624 | -89.0844
8.8627 | -89.0857
8.863 | -89.0867
8.8634 | -89.0877
8.8637 | -89.0883
8.864 | -89.0894
8.8643 | -89.09
8.8647 | -89.0908
8.865 | -89.0929
8.8653 | -89.0942
8.8657 | -89.095
8.866 | -89.0959
8.8663 | -89.0971
8.8666 | -89.0982
8.867 | -89.0991
8.8673 | -89.0998
8.8676 | -89.1005
8.868 | -89.1009
8.8683 | -89.1017
8.8686 | -89.1026
8.8689 | -89.1034
8.8693 | -89.1046
8.8696 | -89.1057
8.8699 | -89.1059
8.8703 | -89.1067
8.8706 | -89.1074
8.8709 | -89.1084
8.8713 | -89.1098
8.8716 | -89.1109
8.8719 | -89.1116
8.8722 | -89.1125
8.8726 | -89.1133
8.8729 | -89.114
8.8732 | -89.1151
8.8736 | -89.1161
8.8739 | -89.1169
8.8742 | -89.1176
8.8745 | -89.1181
8.8749 | -89.1188
8.8752 | -89.1194
8.8755 | -89.1195
8.8759 | -89.1198
8.8762 | -89.1211
8.8765 | -89.1226
8.8768 | -89.1236
8.8772 | -89.1242
8.8775 | -89.1244
8.8778 | -89.126
8.8782 | -89.1268
8.8785 | -89.1279
8.8788 | -89.1286
8.8792 | -89.1291
8.8795 | -89.1297
8.8798 | -89.1304
8.8801 | -89.1311
8.8805 | -89.1322
8.8808 | -89.1329
8.8811 | -89.135
8.8815 | -89.137
8.8818 | -89.1383
8.8821 | -89.14
8.8824 | -89.1409
8.8828 | -89.1418
8.8831 | -89.1417
8.8834 | -89.1422
8.8838 | -89.1428
8.8841 | -89.1436
8.8844 | -89.1444
8.8847 | -89.1455
8.8851 | -89.1468
8.8854 | -89.1474
8.8857 | -89.148
8.8861 | -89.1486
8.8864 | -89.1492
8.8867 | -89.1498
8.8871 | -89.1502
8.8874 | -89.1502
8.8877 | -89.1507
8.888 | -89.1518
8.8884 | -89.1526
8.8887 | -89.1534
8.889 | -89.154
8.8894 | -89.1548
8.8897 | -89.1552
8.89 | -89.1558
8.8903 | -89.1565
8.8907 | -89.1569
8.891 | -89.1572
8.8913 | -89.158
8.8917 | -89.1587
8.892 | -89.1599
8.8923 | -89.1611
8.8926 | -89.1618
8.893 | -89.1626
8.8933 | -89.1638
8.8936 | -89.1648
8.894 | -89.1657
8.8943 | -89.1666
8.8946 | -89.1673
8.895 | -89.168
8.8953 | -89.1687
8.8956 | -89.1693
8.8959 | -89.1696
8.8963 | -89.1697
8.8966 | -89.1707
8.8969 | -89.1716
8.8973 | -89.1722
8.8976 | -89.1732
8.8979 | -89.174
8.8982 | -89.1753
8.8986 | -89.1761
8.8989 | -89.177
8.8992 | -89.1775
8.8996 | -89.1776
8.8999 | -89.1781
8.9002 | -89.1788
8.9005 | -89.1791
8.9009 | -89.1803
8.9012 | -89.1822
8.9015 | -89.183
8.9019 | -89.1835
8.9022 | -89.1841
8.9025 | -89.1849
8.9029 | -89.1858
8.9032 | -89.1857
8.9035 | -89.1854
8.9038 | -89.1861
8.9042 | -89.1862
8.9045 | -89.1868
8.9048 | -89.1877
8.9052 | -89.1884
8.9055 | -89.1887
8.9058 | -89.1888
8.9061 | -89.1895
8.9065 | -89.1903
8.9068 | -89.1909
8.9071 | -89.1917
8.9075 | -89.1912
8.9078 | -89.1915
8.9081 | -89.1921
8.9084 | -89.1924
8.9088 | -89.1933
8.9091 | -89.1947
8.9094 | -89.1952
8.9098 | -89.1963
8.9101 | -89.1973
8.9104 | -89.1978
8.9108 | -89.1986
8.9111 | -89.1991
8.9114 | -89.1998
8.9117 | -89.2005
8.9121 | -89.2012
8.9124 | -89.2018
8.9127 | -89.202
8.9131 | -89.2032
8.9134 | -89.2044
8.9137 | -89.2047
8.914 | -89.2053
8.9144 | -89.2062
8.9147 | -89.2067
8.915 | -89.2078
8.9154 | -89.2082
8.9157 | -89.2093
8.916 | -89.2099
8.9163 | -89.2107
8.9167 | -89.2115
8.917 | -89.2116
8.9173 | -89.212
8.9177 | -89.2127
8.918 | -89.2132
8.9183 | -89.2144
8.9186 | -89.2154
8.919 | -89.2158
8.9193 | -89.2168
8.9196 | -89.2176
8.92 | -89.2175
8.9203 | -89.2177
8.9206 | -89.2182
8.921 | -89.2188
8.9213 | -89.219
8.9216 | -89.22
8.9219 | -89.2206
8.9223 | -89.2212
8.9226 | -89.2218
8.9229 | -89.222
8.9233 | -89.2228
8.9236 | -89.2236
8.9239 | -89.2245
8.9242 | -89.2251
8.9246 | -89.2261
8.9249 | -89.227
8.9252 | -89.2278
8.9256 | -89.2283
8.9259 | -89.2288
8.9262 | -89.2294
8.9266 | -89.2303
8.9269 | -89.2309
8.9272 | -89.2319
8.9275 | -89.2331
8.9279 | -89.2341
8.9282 | -89.2351
8.9285 | -89.2357
8.9289 | -89.2364
8.9292 | -89.2371
8.9295 | -89.2378
8.9298 | -89.2386
8.9302 | -89.2395
8.9305 | -89.2407
8.9308 | -89.2414
8.9312 | -89.2419
8.9315 | -89.2425
8.9318 | -89.2439
8.9321 | -89.2451
8.9325 | -89.2455
8.9328 | -89.2456
8.9331 | -89.2462
8.9335 | -89.2467
8.9338 | -89.2471
8.9341 | -89.2481
8.9344 | -89.2491
8.9348 | -89.2492
8.9351 | -89.2498
8.9354 | -89.2506
8.9358 | -89.2523
8.9361 | -89.2529
8.9364 | -89.2542
8.9368 | -89.2552
8.9371 | -89.2564
8.9374 | -89.2574
8.9377 | -89.2581
8.9381 | -89.2588
8.9384 | -89.2597
8.9387 | -89.2605
8.9391 | -89.2616
8.9394 | -89.2627
8.9397 | -89.2635
8.94 | -89.2641
8.9404 | -89.2646
8.9407 | -89.2656
8.941 | -89.2667
8.9414 | -89.2671
8.9417 | -89.2677
8.942 | -89.268
8.9423 | -89.2688
8.9427 | -89.2696
8.943 | -89.27
8.9433 | -89.2693
8.9437 | -89.2678
8.944 | -89.2675
8.9443 | -89.268
8.9447 | -89.2686
8.945 | -89.27
8.9453 | -89.2706
8.9456 | -89.2712
8.946 | -89.2717
8.9463 | -89.2731
8.9466 | -89.274
8.947 | -89.2744
8.9473 | -89.2746
8.9476 | -89.2751
8.9479 | -89.2763
8.9483 | -89.2787
8.9486 | -89.2811
8.9489 | -89.2823
8.9493 | -89.2828
8.9496 | -89.2832
8.9499 | -89.2837
8.9502 | -89.2839
8.9506 | -89.2844
8.9509 | -89.285
8.9512 | -89.2856
8.9516 | -89.2859
8.9519 | -89.2864
8.9522 | -89.2875
8.9526 | -89.2886
8.9529 | -89.2893
8.9532 | -89.2893
8.9535 | -89.2897
8.9539 | -89.2904
8.9542 | -89.2908
8.9545 | -89.2913
8.9549 | -89.2916
8.9552 | -89.2926
8.9555 | -89.294
8.9558 | -89.2948
8.9562 | -89.2954
8.9565 | -89.2964
8.9568 | -89.2978
8.9572 | -89.299
8.9575 | -89.2996
8.9578 | -89.3002
8.9581 | -89.3006
8.9585 | -89.301
8.9588 | -89.3016
8.9591 | -89.3024
8.9595 | -89.3028
8.9598 | -89.3032
8.9601 | -89.3037
8.9605 | -89.3044
8.9608 | -89.3056
8.9611 | -89.3074
8.9614 | -89.3088
8.9618 | -89.3095
8.9621 | -89.3099
8.9624 | -89.3104
8.9628 | -89.3106
8.9631 | -89.311
8.9634 | -89.3113
8.9637 | -89.3113
8.9641 | -89.3116
8.9644 | -89.3124
8.9647 | -89.313
8.9651 | -89.3141
8.9654 | -89.3145
8.9657 | -89.315
8.966 | -89.3155
8.9664 | -89.316
8.9667 | -89.3162
8.967 | -89.3165
8.9674 | -89.3171
8.9677 | -89.3174
8.968 | -89.318
8.9683 | -89.3188
8.9687 | -89.3197
8.969 | -89.3204
8.9693 | -89.3208
8.9697 | -89.321
8.97 | -89.3213
8.9703 | -89.3213
8.9707 | -89.3219
8.971 | -89.3229
8.9713 | -89.3237
8.9716 | -89.3243
8.972 | -89.3251
8.9723 | -89.3253
8.9726 | -89.3258
8.973 | -89.3263
8.9733 | -89.3268
8.9736 | -89.3272
8.9739 | -89.3276
8.9743 | -89.328
8.9746 | -89.3283
8.9749 | -89.3282
8.9753 | -89.328
8.9756 | -89.3285
8.9759 | -89.3291
8.9762 | -89.3297
8.9766 | -89.33
8.9769 | -89.3304
8.9772 | -89.3315
8.9776 | -89.3322
8.9779 | -89.3328
8.9782 | -89.3334
8.9786 | -89.3339
8.9789 | -89.3345
8.9792 | -89.3354
8.9795 | -89.3364
8.9799 | -89.3369
8.9802 | -89.3376
8.9805 | -89.3382
8.9809 | -89.3387
8.9812 | -89.3391
8.9815 | -89.3392
8.9818 | -89.3392
8.9822 | -89.3397
8.9825 | -89.3401
8.9828 | -89.3401
8.9832 | -89.3404
8.9835 | -89.3411
8.9838 | -89.3419
8.9841 | -89.3422
8.9845 | -89.3421
8.9848 | -89.3419
8.9851 | -89.3419
8.9855 | -89.3423
8.9858 | -89.3426
8.9861 | -89.3438
8.9865 | -89.3451
8.9868 | -89.3453
8.9871 | -89.3454
8.9874 | -89.3459
8.9878 | -89.3465
8.9881 | -89.3473
8.9884 | -89.3477
8.9888 | -89.3479
8.9891 | -89.3481
8.9894 | -89.3488
8.9897 | -89.3486
8.9901 | -89.3487
8.9904 | -89.3491
8.9907 | -89.3499
8.9911 | -89.3503
8.9914 | -89.3508
8.9917 | -89.3512
8.992 | -89.3523
8.9924 | -89.3537
8.9927 | -89.354
8.993 | -89.3548
8.9934 | -89.3551
8.9937 | -89.3558
8.994 | -89.3563
8.9944 | -89.3571
8.9947 | -89.3581
8.995 | -89.3592
8.9953 | -89.3598
8.9957 | -89.3602
8.996 | -89.3604
8.9963 | -89.3607
8.9967 | -89.3612
8.997 | -89.3618
8.9973 | -89.362
8.9976 | -89.3619
8.998 | -89.3623
8.9983 | -89.3625
8.9986 | -89.3625
8.999 | -89.3628
8.9993 | -89.3635
8.9996 | -89.3637
8.9999 | -89.364
9.0003 | -89.3645
9.0006 | -89.3649
9.0009 | -89.3651
9.0013 | -89.365
9.0016 | -89.3647
9.0019 | -89.3652
9.0023 | -89.3663
9.0026 | -89.3663
9.0029 | -89.3663
9.0032 | -89.3667
9.0036 | -89.3675
9.0039 | -89.3681
9.0042 | -89.3684
9.0046 | -89.3688
9.0049 | -89.369
9.0052 | -89.3694
9.0055 | -89.3698
9.0059 | -89.3703
9.0062 | -89.3702
9.0065 | -89.3701
9.0069 | -89.3705
9.0072 | -89.3712
9.0075 | -89.3722
9.0078 | -89.3727
9.0082 | -89.373
9.0085 | -89.3733
9.0088 | -89.3739
9.0092 | -89.3741
9.0095 | -89.3743
9.0098 | -89.3748
9.0102 | -89.3746
9.0105 | -89.3738
9.0108 | -89.3744
9.0111 | -89.375
9.0115 | -89.3747
9.0118 | -89.3745
9.0121 | -89.3747
9.0125 | -89.375
9.0128 | -89.3757
9.0131 | -89.377
9.0134 | -89.378
9.0138 | -89.3785
9.0141 | -89.3788
9.0144 | -89.3792
9.0148 | -89.3796
9.0151 | -89.3798
9.0154 | -89.3802
9.0157 | -89.3803
9.0161 | -89.3816
9.0164 | -89.3828
9.0167 | -89.3832
9.0171 | -89.3834
9.0174 | -89.384
9.0177 | -89.3844
9.0181 | -89.3849
9.0184 | -89.3852
9.0187 | -89.3852
9.019 | -89.3857
9.0194 | -89.3864
9.0197 | -89.3866
9.02 | -89.3869
9.0204 | -89.3873
9.0207 | -89.3878
9.021 | -89.3889
9.0213 | -89.3897
9.0217 | -89.3907
9.022 | -89.3911
9.0223 | -89.3914
9.0227 | -89.392
9.023 | -89.3929
9.0233 | -89.3937
9.0236 | -89.3943
9.024 | -89.3951
9.0243 | -89.3955
9.0246 | -89.3953
9.025 | -89.3957
9.0253 | -89.3962
9.0256 | -89.3964
9.0259 | -89.3964
9.0263 | -89.3969
9.0266 | -89.3972
9.0269 | -89.3977
9.0273 | -89.3982
9.0276 | -89.3989
9.0279 | -89.3994
9.0283 | -89.3996
9.0286 | -89.3995
9.0289 | -89.3994
9.0292 | -89.3997
9.0296 | -89.4
9.0299 | -89.4004
9.0302 | -89.4006
9.0306 | -89.4015
9.0309 | -89.4022
9.0312 | -89.4021
9.0315 | -89.4026
9.0319 | -89.4034
9.0322 | -89.4038
9.0325 | -89.4041
9.0329 | -89.4045
9.0332 | -89.4045
9.0335 | -89.4033
9.0338 | -89.4026
9.0342 | -89.4025
9.0345 | -89.4026
9.0348 | -89.403
9.0352 | -89.4034
9.0355 | -89.4036
9.0358 | -89.4039
9.0362 | -89.4047
9.0365 | -89.4051
9.0368 | -89.4053
9.0371 | -89.4057
9.0375 | -89.4061
9.0378 | -89.4066
9.0381 | -89.4073
9.0385 | -89.4075
9.0388 | -89.4079
9.0391 | -89.4083
9.0394 | -89.4087
9.0398 | -89.4088
9.0401 | -89.4095
9.0404 | -89.41
9.0408 | -89.4113
9.0411 | -89.4117
9.0414 | -89.4119
9.0417 | -89.412
9.0421 | -89.4129
9.0424 | -89.4134
9.0427 | -89.4139
9.0431 | -89.414
9.0434 | -89.4136
9.0437 | -89.4139
9.0441 | -89.4143
9.0444 | -89.415
9.0447 | -89.4155
9.045 | -89.4155
9.0454 | -89.4157
9.0457 | -89.4163
9.046 | -89.4171
9.0464 | -89.4177
9.0467 | -89.4181
9.047 | -89.4187
9.0473 | -89.4189
9.0477 | -89.4192
9.048 | -89.4197
9.0483 | -89.4201
9.0487 | -89.4206
9.049 | -89.421
9.0493 | -89.4213
9.0496 | -89.4213
9.05 | -89.4216
9.0503 | -89.4223
9.0506 | -89.423
9.051 | -89.4235
9.0513 | -89.424
9.0516 | -89.4241
9.052 | -89.425
9.0523 | -89.4255
9.0526 | -89.4256
9.0529 | -89.4257
9.0533 | -89.4261
9.0536 | -89.4263
9.0539 | -89.4248
9.0543 | -89.423
9.0546 | -89.4236
9.0549 | -89.4246
9.0552 | -89.425
9.0556 | -89.4251
9.0559 | -89.4254
9.0562 | -89.4252
9.0566 | -89.425
9.0569 | -89.4252
9.0572 | -89.4254
9.0575 | -89.4254
9.0579 | -89.4256
9.0582 | -89.4262
9.0585 | -89.427
9.0589 | -89.4273
9.0592 | -89.4276
9.0595 | -89.4278
9.0599 | -89.4287
9.0602 | -89.4296
9.0605 | -89.4299
9.0608 | -89.4302
9.0612 | -89.4309
9.0615 | -89.4309
9.0618 | -89.4307
9.0622 | -89.4308
9.0625 | -89.4311
9.0628 | -89.4315
9.0631 | -89.4325
9.0635 | -89.4328
9.0638 | -89.433
9.0641 | -89.4333
9.0645 | -89.4335
9.0648 | -89.4338
9.0651 | -89.434
9.0654 | -89.4352
9.0658 | -89.4358
9.0661 | -89.4363
9.0664 | -89.4375
9.0668 | -89.4383
9.0671 | -89.4388
9.0674 | -89.4396
9.0678 | -89.4397
9.0681 | -89.4403
9.0684 | -89.441
9.0687 | -89.4412
9.0691 | -89.4417
9.0694 | -89.4423
9.0697 | -89.4425
9.0701 | -89.4425
9.0704 | -89.4438
9.0707 | -89.4443
9.071 | -89.4445
9.0714 | -89.445
9.0717 | -89.4464
9.072 | -89.4473
9.0724 | -89.4473
9.0727 | -89.4477
9.073 | -89.4481
9.0733 | -89.4484
9.0737 | -89.4487
9.074 | -89.4491
9.0743 | -89.4499
9.0747 | -89.4506
9.075 | -89.4509
9.0753 | -89.4516
9.0756 | -89.4522
9.076 | -89.4524
9.0763 | -89.4531
9.0766 | -89.4538
9.077 | -89.4541
9.0773 | -89.4546
9.0776 | -89.4553
9.078 | -89.4555
9.0783 | -89.4561
9.0786 | -89.4565
9.0789 | -89.4571
9.0793 | -89.458
9.0796 | -89.4586
9.0799 | -89.4595
9.0803 | -89.4598
9.0806 | -89.4604
9.0809 | -89.4611
9.0812 | -89.4626
9.0816 | -89.4642
9.0819 | -89.4641
9.0822 | -89.4648
9.0826 | -89.4653
9.0829 | -89.4669
9.0832 | -89.4692
9.0835 | -89.4703
9.0839 | -89.4703
9.0842 | -89.4707
9.0845 | -89.4714
9.0849 | -89.4719
9.0852 | -89.4724
9.0855 | -89.4728
9.0859 | -89.4731
9.0862 | -89.4734
9.0865 | -89.4744
9.0868 | -89.475
9.0872 | -89.475
9.0875 | -89.4754
9.0878 | -89.475
9.0882 | -89.4757
9.0885 | -89.4775
9.0888 | -89.4785
9.0891 | -89.4788
9.0895 | -89.4796
9.0898 | -89.4807
9.0901 | -89.4814
9.0905 | -89.482
9.0908 | -89.4825
9.0911 | -89.483
9.0914 | -89.4828
9.0918 | -89.4832
9.0921 | -89.4834
9.0924 | -89.4836
9.0928 | -89.4839
9.0931 | -89.4847
9.0934 | -89.4848
9.0938 | -89.485
9.0941 | -89.4853
9.0944 | -89.4854
9.0947 | -89.486
9.0951 | -89.4865
9.0954 | -89.4866
9.0957 | -89.4873
9.0961 | -89.4879
9.0964 | -89.4882
9.0967 | -89.4882
9.097 | -89.4877
9.0974 | -89.488
9.0977 | -89.4883
9.098 | -89.4885
9.0984 | -89.4891
9.0987 | -89.4897
9.099 | -89.4905
9.0993 | -89.4905
9.0997 | -89.491
9.1 | -89.4914
9.1003 | -89.4915
9.1007 | -89.4915
9.101 | -89.4918
9.1013 | -89.4921
9.1017 | -89.4925
9.102 | -89.4925
9.1023 | -89.4926
9.1026 | -89.4928
9.103 | -89.4933
9.1033 | -89.4938
9.1036 | -89.494
9.104 | -89.4943
9.1043 | -89.4946
9.1046 | -89.4947
9.1049 | -89.495
9.1053 | -89.4946
9.1056 | -89.494
9.1059 | -89.4945
9.1063 | -89.4943
9.1066 | -89.4933
9.1069 | -89.4936
9.1072 | -89.4939
9.1076 | -89.4941
9.1079 | -89.4945
9.1082 | -89.4948
9.1086 | -89.4952
9.1089 | -89.4952
9.1092 | -89.4945
9.1096 | -89.4938
9.1099 | -89.4936
9.1102 | -89.4938
9.1105 | -89.4937
9.1109 | -89.4936
9.1112 | -89.4939
9.1115 | -89.4944
9.1119 | -89.495
9.1122 | -89.4951
9.1125 | -89.4952
9.1128 | -89.4953
9.1132 | -89.4957
9.1135 | -89.4961
9.1138 | -89.4962
9.1142 | -89.4967
9.1145 | -89.4975
9.1148 | -89.4975
9.1151 | -89.4976
9.1155 | -89.4978
9.1158 | -89.4982
9.1161 | -89.4987
9.1165 | -89.4991
9.1168 | -89.5002
9.1171 | -89.5008
9.1174 | -89.5005
9.1178 | -89.5003
9.1181 | -89.5005
9.1184 | -89.5006
9.1188 | -89.5012
9.1191 | -89.5014
9.1194 | -89.5017
9.1198 | -89.5022
9.1201 | -89.5025
9.1204 | -89.5031
9.1207 | -89.5036
9.1211 | -89.5041
9.1214 | -89.5046
9.1217 | -89.5051
9.1221 | -89.5057
9.1224 | -89.5064
9.1227 | -89.5072
9.123 | -89.5077
9.1234 | -89.5081
9.1237 | -89.5084
9.124 | -89.5087
9.1244 | -89.5088
9.1247 | -89.5094
9.125 | -89.5095
9.1253 | -89.5097
9.1257 | -89.5102
9.126 | -89.5109
9.1263 | -89.5112
9.1267 | -89.5112
9.127 | -89.5118
9.1273 | -89.5123
9.1277 | -89.5125
9.128 | -89.5129
9.1283 | -89.5131
9.1286 | -89.5135
9.129 | -89.5142
9.1293 | -89.515
9.1296 | -89.5162
9.13 | -89.5165
9.1303 | -89.5169
9.1306 | -89.517
9.1309 | -89.5178
9.1313 | -89.5186
9.1316 | -89.519
9.1319 | -89.5202
9.1323 | -89.5208
9.1326 | -89.5213
9.1329 | -89.522
9.1332 | -89.5229
9.1336 | -89.5236
9.1339 | -89.524
9.1342 | -89.5239
9.1346 | -89.5241
9.1349 | -89.5241
9.1352 | -89.5244
9.1356 | -89.5239
9.1359 | -89.5229
9.1362 | -89.5225
9.1365 | -89.5226
9.1369 | -89.5231
9.1372 | -89.524
9.1375 | -89.5242
9.1379 | -89.5252
9.1382 | -89.5255
9.1385 | -89.5253
9.1388 | -89.5253
9.1392 | -89.5253
9.1395 | -89.5248
9.1398 | -89.5255
9.1402 | -89.526
9.1405 | -89.5263
9.1408 | -89.5266
9.1411 | -89.5268
9.1415 | -89.5271
9.1418 | -89.5274
9.1421 | -89.5282
9.1425 | -89.5294
9.1428 | -89.5311
9.1431 | -89.5317
9.1435 | -89.532
9.1438 | -89.5321
9.1441 | -89.5323
9.1444 | -89.5326
9.1448 | -89.5332
9.1451 | -89.5336
9.1454 | -89.5339
9.1458 | -89.5346
9.1461 | -89.535
9.1464 | -89.5352
9.1467 | -89.5356
9.1471 | -89.536
9.1474 | -89.5365
9.1477 | -89.5368
9.1481 | -89.537
9.1484 | -89.5372
9.1487 | -89.5379
9.149 | -89.5379
9.1494 | -89.5384
9.1497 | -89.5388
9.15 | -89.5391
9.1504 | -89.5393
9.1507 | -89.5393
9.151 | -89.5398
9.1514 | -89.54
9.1517 | -89.5404
9.152 | -89.5407
9.1523 | -89.5412
9.1527 | -89.5417
9.153 | -89.5424
9.1533 | -89.5427
9.1537 | -89.5427
9.154 | -89.5428
9.1543 | -89.5426
9.1546 | -89.5428
9.155 | -89.5433
9.1553 | -89.5441
9.1556 | -89.5448
9.156 | -89.5451
9.1563 | -89.5449
9.1566 | -89.5441
9.1569 | -89.5437
9.1573 | -89.5438
9.1576 | -89.5437
9.1579 | -89.5435
9.1583 | -89.5436
9.1586 | -89.544
9.1589 | -89.5443
9.1593 | -89.5449
9.1596 | -89.5455
9.1599 | -89.5455
9.1602 | -89.5459
9.1606 | -89.5459
9.1609 | -89.546
9.1612 | -89.5464
9.1616 | -89.547
9.1619 | -89.5473
9.1622 | -89.5474
9.1625 | -89.5475
9.1629 | -89.548
9.1632 | -89.5485
9.1635 | -89.549
9.1639 | -89.5496
9.1642 | -89.5502
9.1645 | -89.5511
9.1648 | -89.5526
9.1652 | -89.5532
9.1655 | -89.553
9.1658 | -89.5529
9.1662 | -89.554
9.1665 | -89.5546
9.1668 | -89.5547
9.1671 | -89.5551
9.1675 | -89.5555
9.1678 | -89.5556
9.1681 | -89.5558
9.1685 | -89.5568
9.1688 | -89.558
9.1691 | -89.5587
9.1695 | -89.559
9.1698 | -89.5593
9.1701 | -89.5591
9.1704 | -89.5594
9.1708 | -89.5596
9.1711 | -89.5604
9.1714 | -89.5606
9.1718 | -89.5611
9.1721 | -89.5624
9.1724 | -89.5628
9.1727 | -89.5632
9.1731 | -89.5633
9.1734 | -89.5639
9.1737 | -89.5643
9.1741 | -89.5647
9.1744 | -89.565
9.1747 | -89.5653
9.175 | -89.5658
9.1754 | -89.5665
9.1757 | -89.5668
9.176 | -89.5672
9.1764 | -89.5681
9.1767 | -89.5684
9.177 | -89.5689
9.1774 | -89.5694
9.1777 | -89.5694
9.178 | -89.5697
9.1783 | -89.5701
9.1787 | -89.5706
9.179 | -89.5712
9.1793 | -89.5715
9.1797 | -89.5718
9.18 | -89.5727
9.1803 | -89.5743
9.1806 | -89.5757
9.181 | -89.5769
9.1813 | -89.5771
9.1816 | -89.5772
9.182 | -89.5766
9.1823 | -89.5766
9.1826 | -89.577
9.1829 | -89.5775
9.1833 | -89.578
9.1836 | -89.5783
9.1839 | -89.5785
9.1843 | -89.579
9.1846 | -89.5795
9.1849 | -89.58
9.1853 | -89.5804
9.1856 | -89.5811
9.1859 | -89.5813
9.1862 | -89.5808
9.1866 | -89.5809
9.1869 | -89.5815
9.1872 | -89.5825
9.1876 | -89.5834
9.1879 | -89.5837
9.1882 | -89.5841
9.1885 | -89.5848
9.1889 | -89.5845
9.1892 | -89.5846
9.1895 | -89.5848
9.1899 | -89.5849
9.1902 | -89.5853
9.1905 | -89.5858
9.1908 | -89.5864
9.1912 | -89.5869
9.1915 | -89.587
9.1918 | -89.5866
9.1922 | -89.5865
9.1925 | -89.5871
9.1928 | -89.5871
9.1932 | -89.5873
9.1935 | -89.5878
9.1938 | -89.5885
9.1941 | -89.5886
9.1945 | -89.5889
9.1948 | -89.5893
9.1951 | -89.5892
9.1955 | -89.5895
9.1958 | -89.5894
9.1961 | -89.5891
9.1964 | -89.5895
9.1968 | -89.5908
9.1971 | -89.5909
9.1974 | -89.5914
9.1978 | -89.5916
9.1981 | -89.5921
9.1984 | -89.5926
9.1987 | -89.593
9.1991 | -89.5933
9.1994 | -89.5934
9.1997 | -89.5935
9.2001 | -89.5937
9.2004 | -89.5943
9.2007 | -89.5946
9.2011 | -89.5954
9.2014 | -89.5963
9.2017 | -89.5968
9.202 | -89.5972
9.2024 | -89.5974
9.2027 | -89.5974
9.203 | -89.5977
9.2034 | -89.598
9.2037 | -89.5984
9.204 | -89.5987
9.2043 | -89.5992
9.2047 | -89.5995
9.205 | -89.5998
9.2053 | -89.6
9.2057 | -89.6002
9.206 | -89.6005
9.2063 | -89.6
9.2066 | -89.5997
9.207 | -89.6
9.2073 | -89.6001
9.2076 | -89.6005
9.208 | -89.6008
9.2083 | -89.6011
9.2086 | -89.6018
9.2089 | -89.6031
9.2093 | -89.6036
9.2096 | -89.6043
9.2099 | -89.6049
9.2103 | -89.6042
9.2106 | -89.6047
9.2109 | -89.6053
9.2113 | -89.6056
9.2116 | -89.6057
9.2119 | -89.6059
9.2122 | -89.6062
9.2126 | -89.6066
9.2129 | -89.6064
9.2132 | -89.6064
9.2136 | -89.6066
9.2139 | -89.6066
9.2142 | -89.6069
9.2145 | -89.6072
9.2149 | -89.6074
9.2152 | -89.608
9.2155 | -89.6085
9.2159 | -89.6084
9.2162 | -89.6081
9.2165 | -89.6082
9.2168 | -89.6093
9.2172 | -89.6089
9.2175 | -89.6076
9.2178 | -89.6075
9.2182 | -89.608
9.2185 | -89.6076
9.2188 | -89.6077
9.2192 | -89.6078
9.2195 | -89.6079
9.2198 | -89.6079
9.2201 | -89.6074
9.2205 | -89.607
9.2208 | -89.6071
9.2211 | -89.6073
9.2215 | -89.6072
9.2218 | -89.6073
9.2221 | -89.6081
9.2224 | -89.6089
9.2228 | -89.6097
9.2231 | -89.6098
9.2234 | -89.6099
9.2238 | -89.61
9.2241 | -89.6106
9.2244 | -89.6109
9.2247 | -89.6111
9.2251 | -89.6114
9.2254 | -89.6114
9.2257 | -89.6114
9.2261 | -89.6115
9.2264 | -89.6116
9.2267 | -89.612
9.2271 | -89.6122
9.2274 | -89.6124
9.2277 | -89.6126
9.228 | -89.6127
9.2284 | -89.6129
9.2287 | -89.6131
9.229 | -89.6135
9.2294 | -89.6137
9.2297 | -89.6133
9.23 | -89.6135
9.2303 | -89.6141
9.2307 | -89.6147
9.231 | -89.6152
9.2313 | -89.6154
9.2317 | -89.6154
9.232 | -89.6155
9.2323 | -89.6164
9.2326 | -89.6179
9.233 | -89.6183
9.2333 | -89.6181
9.2336 | -89.6185
9.234 | -89.6188
9.2343 | -89.6188
9.2346 | -89.619
9.235 | -89.6194
9.2353 | -89.62
9.2356 | -89.6204
9.2359 | -89.6206
9.2363 | -89.6211
9.2366 | -89.6211
9.2369 | -89.6215
9.2373 | -89.6214
9.2376 | -89.6214
9.2379 | -89.622
9.2382 | -89.6232
9.2386 | -89.6241
9.2389 | -89.6242
9.2392 | -89.6243
9.2396 | -89.6248
9.2399 | -89.625
9.2402 | -89.6253
9.2405 | -89.6255
9.2409 | -89.6258
9.2412 | -89.6259
9.2415 | -89.6259
9.2419 | -89.6265
9.2422 | -89.6269
9.2425 | -89.6275
9.2429 | -89.6279
9.2432 | -89.6282
9.2435 | -89.6286
9.2438 | -89.6289
9.2442 | -89.629
9.2445 | -89.6293
9.2448 | -89.6292
9.2452 | -89.6295
9.2455 | -89.63
9.2458 | -89.6303
9.2461 | -89.6307
9.2465 | -89.6307
9.2468 | -89.6306
9.2471 | -89.6305
9.2475 | -89.6303
9.2478 | -89.6304
9.2481 | -89.6306
9.2484 | -89.6307
9.2488 | -89.6308
9.2491 | -89.6309
9.2494 | -89.6315
9.2498 | -89.6318
9.2501 | -89.6321
9.2504 | -89.6326
9.2508 | -89.6332
9.2511 | -89.6338
9.2514 | -89.6341
9.2517 | -89.6341
9.2521 | -89.6336
9.2524 | -89.6336
9.2527 | -89.6336
9.2531 | -89.6338
9.2534 | -89.6344
9.2537 | -89.6351
9.254 | -89.6355
9.2544 | -89.6355
9.2547 | -89.6358
9.255 | -89.636
9.2554 | -89.6363
9.2557 | -89.6366
9.256 | -89.6369
9.2563 | -89.6371
9.2567 | -89.6379
9.257 | -89.639
9.2573 | -89.6392
9.2577 | -89.6392
9.258 | -89.6397
9.2583 | -89.6398
9.2586 | -89.6396
9.259 | -89.6399
9.2593 | -89.6402
9.2596 | -89.6402
9.26 | -89.64
9.2603 | -89.6402
9.2606 | -89.6404
9.261 | -89.6407
9.2613 | -89.6409
9.2616 | -89.6412
9.2619 | -89.6412
9.2623 | -89.6411
9.2626 | -89.6409
9.2629 | -89.6409
9.2633 | -89.6414
9.2636 | -89.6419
9.2639 | -89.6424
9.2642 | -89.6427
9.2646 | -89.6427
9.2649 | -89.6432
9.2652 | -89.644
9.2656 | -89.6441
9.2659 | -89.6438
9.2662 | -89.644
9.2665 | -89.6442
9.2669 | -89.6445
9.2672 | -89.6445
9.2675 | -89.645
9.2679 | -89.6453
9.2682 | -89.6454
9.2685 | -89.6458
9.2689 | -89.6458
9.2692 | -89.6451
9.2695 | -89.6447
9.2698 | -89.6448
9.2702 | -89.6451
9.2705 | -89.6453
9.2708 | -89.6455
9.2712 | -89.6457
9.2715 | -89.6459
9.2718 | -89.6463
9.2721 | -89.6465
9.2725 | -89.6465
9.2728 | -89.647
9.2731 | -89.6472
9.2735 | -89.6475
9.2738 | -89.6477
9.2741 | -89.648
9.2744 | -89.6482
9.2748 | -89.6486
9.2751 | -89.6489
9.2754 | -89.6492
9.2758 | -89.6494
9.2761 | -89.6495
9.2764 | -89.6498
9.2768 | -89.65
9.2771 | -89.6503
9.2774 | -89.6506
9.2777 | -89.6506
9.2781 | -89.6507
9.2784 | -89.6508
9.2787 | -89.6511
9.2791 | -89.6516
9.2794 | -89.6522
9.2797 | -89.6527
9.28 | -89.6527
9.2804 | -89.653
9.2807 | -89.6531
9.281 | -89.6534
9.2814 | -89.6532
9.2817 | -89.6537
9.282 | -89.6541
9.2823 | -89.6548
9.2827 | -89.6554
9.283 | -89.6557
9.2833 | -89.6566
9.2837 | -89.657
9.284 | -89.6574
9.2843 | -89.6577
9.2847 | -89.6577
9.285 | -89.658
9.2853 | -89.6587
9.2856 | -89.6589
9.286 | -89.6595
9.2863 | -89.6602
9.2866 | -89.6613
9.287 | -89.6618
9.2873 | -89.6615
9.2876 | -89.6619
9.2879 | -89.6621
9.2883 | -89.662
9.2886 | -89.6613
9.2889 | -89.6599
9.2893 | -89.6593
9.2896 | -89.6595
9.2899 | -89.6596
9.2902 | -89.66
9.2906 | -89.6606
9.2909 | -89.661
9.2912 | -89.6622
9.2916 | -89.6629
9.2919 | -89.6636
9.2922 | -89.664
9.2926 | -89.6643
9.2929 | -89.6645
9.2932 | -89.6645
9.2935 | -89.6644
9.2939 | -89.6641
9.2942 | -89.6642
9.2945 | -89.6644
9.2949 | -89.6647
9.2952 | -89.6659
9.2955 | -89.6663
9.2958 | -89.6665
9.2962 | -89.6667
9.2965 | -89.6666
9.2968 | -89.6669
9.2972 | -89.667
9.2975 | -89.6678
9.2978 | -89.6679
9.2981 | -89.6676
9.2985 | -89.6673
9.2988 | -89.6674
9.2991 | -89.668
9.2995 | -89.6683
9.2998 | -89.6681
9.3001 | -89.668
9.3004 | -89.6681
9.3008 | -89.6683
9.3011 | -89.6684
9.3014 | -89.668
9.3018 | -89.6685
9.3021 | -89.6694
9.3024 | -89.6695
9.3028 | -89.6701
9.3031 | -89.6705
9.3034 | -89.6705
9.3037 | -89.6708
9.3041 | -89.6713
9.3044 | -89.6711
9.3047 | -89.6714
9.3051 | -89.6716
9.3054 | -89.6718
9.3057 | -89.6723
9.306 | -89.6731
9.3064 | -89.6726
9.3067 | -89.6724
9.307 | -89.6726
9.3074 | -89.6729
9.3077 | -89.6733
9.308 | -89.6733
9.3084 | -89.6734
9.3087 | -89.6734
9.309 | -89.6732
9.3093 | -89.6734
9.3097 | -89.6726
9.31 | -89.6723
9.3103 | -89.6725
9.3107 | -89.6729
9.311 | -89.6735
9.3113 | -89.6741
9.3116 | -89.6745
9.312 | -89.6748
9.3123 | -89.6746
9.3126 | -89.6742
9.313 | -89.6739
9.3133 | -89.6741
9.3136 | -89.6743
9.3139 | -89.6745
9.3143 | -89.6751
9.3146 | -89.6754
9.3149 | -89.6757
9.3153 | -89.676
9.3156 | -89.676
9.3159 | -89.6765
9.3162 | -89.6767
9.3166 | -89.6768
9.3169 | -89.6771
9.3172 | -89.6776
9.3176 | -89.6785
9.3179 | -89.6788
9.3182 | -89.679
9.3186 | -89.68
9.3189 | -89.6808
9.3192 | -89.6805
9.3195 | -89.6799
9.3199 | -89.6798
9.3202 | -89.6799
9.3205 | -89.68
9.3209 | -89.6802
9.3212 | -89.6813
9.3215 | -89.6822
9.3218 | -89.6825
9.3222 | -89.6832
9.3225 | -89.6833
9.3228 | -89.6835
9.3232 | -89.6837
9.3235 | -89.6841
9.3238 | -89.6838
9.3241 | -89.6837
9.3245 | -89.6835
9.3248 | -89.6834
9.3251 | -89.6835
9.3255 | -89.6839
9.3258 | -89.6839
9.3261 | -89.6841
9.3265 | -89.6843
9.3268 | -89.6848
9.3271 | -89.6851
9.3274 | -89.6855
9.3278 | -89.6864
9.3281 | -89.6868
9.3284 | -89.6872
9.3288 | -89.6877
9.3291 | -89.6878
9.3294 | -89.6876
9.3297 | -89.6872
9.3301 | -89.687
9.3304 | -89.6872
9.3307 | -89.6872
9.3311 | -89.6874
9.3314 | -89.6885
9.3317 | -89.6889
9.332 | -89.6898
9.3324 | -89.6899
9.3327 | -89.6898
9.333 | -89.6899
9.3334 | -89.6901
9.3337 | -89.6903
9.334 | -89.6905
9.3344 | -89.6914
9.3347 | -89.692
9.335 | -89.6927
9.3353 | -89.6931
9.3357 | -89.6936
9.336 | -89.6937
9.3363 | -89.6933
9.3367 | -89.6932
9.337 | -89.6937
9.3373 | -89.6936
9.3376 | -89.6938
9.338 | -89.6934
9.3383 | -89.6938
9.3386 | -89.695
9.339 | -89.6955
9.3393 | -89.6957
9.3396 | -89.6957
9.3399 | -89.6959
9.3403 | -89.6958
9.3406 | -89.6956
9.3409 | -89.6955
9.3413 | -89.6954
9.3416 | -89.6954
9.3419 | -89.6956
9.3423 | -89.6958
9.3426 | -89.6959
9.3429 | -89.6962
9.3432 | -89.6963
9.3436 | -89.6966
9.3439 | -89.6968
9.3442 | -89.6975
9.3446 | -89.6977
9.3449 | -89.6978
9.3452 | -89.6977
9.3455 | -89.6976
9.3459 | -89.698
9.3462 | -89.6983
9.3465 | -89.6984
9.3469 | -89.6985
9.3472 | -89.698
9.3475 | -89.6981
9.3478 | -89.6983
9.3482 | -89.6984
9.3485 | -89.6989
9.3488 | -89.6992
9.3492 | -89.6994
9.3495 | -89.7001
9.3498 | -89.7006
9.3501 | -89.7008
9.3505 | -89.701
9.3508 | -89.7012
9.3511 | -89.7013
9.3515 | -89.7018
9.3518 | -89.7026
9.3521 | -89.7029
9.3525 | -89.7031
9.3528 | -89.7033
9.3531 | -89.7037
9.3534 | -89.7043
9.3538 | -89.7048
9.3541 | -89.7049
9.3544 | -89.7048
9.3548 | -89.7048
9.3551 | -89.7051
9.3554 | -89.7052
9.3557 | -89.7053
9.3561 | -89.705
9.3564 | -89.7053
9.3567 | -89.7054
9.3571 | -89.7057
9.3574 | -89.7058
9.3577 | -89.706
9.358 | -89.706
9.3584 | -89.7061
9.3587 | -89.7063
9.359 | -89.7067
9.3594 | -89.7068
9.3597 | -89.7068
9.36 | -89.707
9.3604 | -89.7074
9.3607 | -89.7084
9.361 | -89.7093
9.3613 | -89.7097
9.3617 | -89.7099
9.362 | -89.7108
9.3623 | -89.7111
9.3627 | -89.7114
9.363 | -89.7117
9.3633 | -89.712
9.3636 | -89.7125
9.364 | -89.7128
9.3643 | -89.713
9.3646 | -89.7133
9.365 | -89.7134
9.3653 | -89.7132
9.3656 | -89.7131
9.3659 | -89.7138
9.3663 | -89.7142
9.3666 | -89.7145
9.3669 | -89.7149
9.3673 | -89.7153
9.3676 | -89.7152
9.3679 | -89.7149
9.3683 | -89.7158
9.3686 | -89.7169
9.3689 | -89.7168
9.3692 | -89.7163
9.3696 | -89.7161
9.3699 | -89.7163
9.3702 | -89.7174
9.3706 | -89.7186
9.3709 | -89.7189
9.3712 | -89.7194
9.3715 | -89.72
9.3719 | -89.7198
9.3722 | -89.7199
9.3725 | -89.7203
9.3729 | -89.721
9.3732 | -89.7215
9.3735 | -89.7216
9.3738 | -89.7216
9.3742 | -89.7219
9.3745 | -89.7219
9.3748 | -89.7218
9.3752 | -89.7217
9.3755 | -89.7219
9.3758 | -89.7221
9.3762 | -89.7224
9.3765 | -89.723
9.3768 | -89.7227
9.3771 | -89.7225
9.3775 | -89.7233
9.3778 | -89.7237
9.3781 | -89.7238
9.3785 | -89.7241
9.3788 | -89.7243
9.3791 | -89.7244
9.3794 | -89.7244
9.3798 | -89.7244
9.3801 | -89.7243
9.3804 | -89.7243
9.3808 | -89.7238
9.3811 | -89.7241
9.3814 | -89.7244
9.3817 | -89.7249
9.3821 | -89.7254
9.3824 | -89.726
9.3827 | -89.7262
9.3831 | -89.7257
9.3834 | -89.7254
9.3837 | -89.7256
9.3841 | -89.7259
9.3844 | -89.7263
9.3847 | -89.7268
9.385 | -89.7276
9.3854 | -89.7284
9.3857 | -89.7287
9.386 | -89.7289
9.3864 | -89.7292
9.3867 | -89.7295
9.387 | -89.7295
9.3873 | -89.7299
9.3877 | -89.7304
9.388 | -89.7309
9.3883 | -89.7315
9.3887 | -89.7319
9.389 | -89.7323
9.3893 | -89.7323
9.3896 | -89.7324
9.39 | -89.7328
9.3903 | -89.7335
9.3906 | -89.7338
9.391 | -89.734
9.3913 | -89.734
9.3916 | -89.734
9.392 | -89.7334
9.3923 | -89.7333
9.3926 | -89.7332
9.3929 | -89.733
9.3933 | -89.7331
9.3936 | -89.7333
9.3939 | -89.7334
9.3943 | -89.7334
9.3946 | -89.7336
9.3949 | -89.7334
9.3952 | -89.7333
9.3956 | -89.7336
9.3959 | -89.7336
9.3962 | -89.7345
9.3966 | -89.7353
9.3969 | -89.7362
9.3972 | -89.7361
9.3975 | -89.7354
9.3979 | -89.7353
9.3982 | -89.7355
9.3985 | -89.7355
9.3989 | -89.7358
9.3992 | -89.7364
9.3995 | -89.7367
9.3999 | -89.7372
9.4002 | -89.7373
9.4005 | -89.7373
9.4008 | -89.7376
9.4012 | -89.7379
9.4015 | -89.7385
9.4018 | -89.7385
9.4022 | -89.7388
9.4025 | -89.7393
9.4028 | -89.7394
9.4031 | -89.7395
9.4035 | -89.7402
9.4038 | -89.7415
9.4041 | -89.7425
9.4045 | -89.7427
9.4048 | -89.7426
9.4051 | -89.7423
9.4054 | -89.743
9.4058 | -89.7433
9.4061 | -89.7433
9.4064 | -89.7435
9.4068 | -89.7437
9.4071 | -89.744
9.4074 | -89.744
9.4077 | -89.744
9.4081 | -89.7438
9.4084 | -89.7443
9.4087 | -89.7446
9.4091 | -89.7446
9.4094 | -89.7451
9.4097 | -89.7449
9.4101 | -89.7449
9.4104 | -89.745
9.4107 | -89.7455
9.411 | -89.7459
9.4114 | -89.7457
9.4117 | -89.7459
9.412 | -89.7461
9.4124 | -89.7464
9.4127 | -89.7464
9.413 | -89.7464
9.4133 | -89.7469
9.4137 | -89.7478
9.414 | -89.7482
9.4143 | -89.7485
9.4147 | -89.7496
9.415 | -89.7498
9.4153 | -89.7502
9.4156 | -89.7504
9.416 | -89.7507
9.4163 | -89.751
9.4166 | -89.7514
9.417 | -89.752
9.4173 | -89.7524
9.4176 | -89.7526
9.418 | -89.7529
9.4183 | -89.7535
9.4186 | -89.754
9.4189 | -89.7539
9.4193 | -89.7552
9.4196 | -89.7543
9.4199 | -89.7548
9.4203 | -89.7555
9.4206 | -89.7561
9.4209 | -89.7564
9.4212 | -89.7577
9.4216 | -89.7585
9.4219 | -89.7587
9.4222 | -89.7589
9.4226 | -89.7596
9.4229 | -89.7598
9.4232 | -89.76
9.4235 | -89.7603
9.4239 | -89.7607
9.4242 | -89.7609
9.4245 | -89.7611
9.4249 | -89.7613
9.4252 | -89.7618
9.4255 | -89.762
9.4259 | -89.7623
9.4262 | -89.7626
9.4265 | -89.7629
9.4268 | -89.7634
9.4272 | -89.7636
9.4275 | -89.7638
9.4278 | -89.7641
9.4282 | -89.7647
9.4285 | -89.7655
9.4288 | -89.7661
9.4291 | -89.7666
9.4295 | -89.7668
9.4298 | -89.7669
9.4301 | -89.7674
9.4305 | -89.7677
9.4308 | -89.7678
9.4311 | -89.7681
9.4314 | -89.7681
9.4318 | -89.7681
9.4321 | -89.7686
9.4324 | -89.7688
9.4328 | -89.7694
9.4331 | -89.7702
9.4334 | -89.7714
9.4338 | -89.7726
9.4341 | -89.7733
9.4344 | -89.7738
9.4347 | -89.7745
9.4351 | -89.7755
9.4354 | -89.7773
9.4357 | -89.7781
9.4361 | -89.7785
9.4364 | -89.7789
9.4367 | -89.7797
9.437 | -89.7801
9.4374 | -89.7804
9.4377 | -89.7805
9.438 | -89.7811
9.4384 | -89.7816
9.4387 | -89.7816
9.439 | -89.7819
9.4393 | -89.7826
9.4397 | -89.7831
9.44 | -89.7832
9.4403 | -89.7834
9.4407 | -89.7837
9.441 | -89.7842
9.4413 | -89.7843
9.4416 | -89.7838
9.442 | -89.783
9.4423 | -89.7829
9.4426 | -89.7836
9.443 | -89.7843
9.4433 | -89.7846
9.4436 | -89.7849
9.444 | -89.7853
9.4443 | -89.7856
9.4446 | -89.7859
9.4449 | -89.7859
9.4453 | -89.7864
9.4456 | -89.7865
9.4459 | -89.7866
9.4463 | -89.7871
9.4466 | -89.7873
9.4469 | -89.7879
9.4472 | -89.7887
9.4476 | -89.7891
9.4479 | -89.7905
9.4482 | -89.7912
9.4486 | -89.7915
9.4489 | -89.7917
9.4492 | -89.7921
9.4496 | -89.7915
9.4499 | -89.7911
9.4502 | -89.7914
9.4505 | -89.7918
9.4509 | -89.7922
9.4512 | -89.7924
9.4515 | -89.7924
9.4519 | -89.793
9.4522 | -89.7937
9.4525 | -89.7936
9.4528 | -89.793
9.4532 | -89.7931
9.4535 | -89.7939
9.4538 | -89.7943
9.4542 | -89.7948
9.4545 | -89.7951
9.4548 | -89.7955
9.4551 | -89.796
9.4555 | -89.7958
9.4558 | -89.7959
9.4561 | -89.7961
9.4565 | -89.7962
9.4568 | -89.7964
9.4571 | -89.7967
9.4574 | -89.7969
9.4578 | -89.7973
9.4581 | -89.7976
9.4584 | -89.7981
9.4588 | -89.7983
9.4591 | -89.7986
9.4594 | -89.7991
9.4598 | -89.7994
9.4601 | -89.7996
9.4604 | -89.7993
9.4607 | -89.7997
9.4611 | -89.8005
9.4614 | -89.8012
9.4617 | -89.8015
9.4621 | -89.802
9.4624 | -89.8023
9.4627 | -89.802
9.463 | -89.8024
9.4634 | -89.804
9.4637 | -89.8068
9.464 | -89.808
9.4644 | -89.8084
9.4647 | -89.8094
9.465 | -89.8106
9.4653 | -89.8108
9.4657 | -89.8109
9.466 | -89.8117
9.4663 | -89.8124
9.4667 | -89.813
9.467 | -89.8136
9.4673 | -89.8139
9.4677 | -89.8143
9.468 | -89.8146
9.4683 | -89.8147
9.4686 | -89.8145
9.469 | -89.8149
9.4693 | -89.8152
9.4696 | -89.8157
9.47 | -89.8165
9.4703 | -89.8168
9.4706 | -89.8173
9.4709 | -89.8173
9.4713 | -89.8173
9.4716 | -89.8179
9.4719 | -89.8183
9.4723 | -89.8187
9.4726 | -89.8194
9.4729 | -89.8201
9.4732 | -89.8206
9.4736 | -89.8207
9.4739 | -89.8213
9.4742 | -89.8216
9.4746 | -89.8221
9.4749 | -89.8226
9.4752 | -89.8223
9.4756 | -89.8225
9.4759 | -89.8228
9.4762 | -89.8229
9.4765 | -89.8231
9.4769 | -89.8233
9.4772 | -89.824
9.4775 | -89.8244
9.4779 | -89.8245
9.4782 | -89.8247
9.4785 | -89.8251
9.4788 | -89.8255
9.4792 | -89.8262
9.4795 | -89.8268
9.4798 | -89.8269
9.4802 | -89.8272
9.4805 | -89.827
9.4808 | -89.8282
9.4811 | -89.8309
9.4815 | -89.8328
9.4818 | -89.8332
9.4821 | -89.8334
9.4825 | -89.8338
9.4828 | -89.834
9.4831 | -89.8345
9.4835 | -89.8346
9.4838 | -89.8337
9.4841 | -89.8338
9.4844 | -89.8341
9.4848 | -89.8343
9.4851 | -89.8346
9.4854 | -89.8349
9.4858 | -89.8352
9.4861 | -89.8358
9.4864 | -89.837
9.4867 | -89.8378
9.4871 | -89.8386
9.4874 | -89.8391
9.4877 | -89.8393
9.4881 | -89.8397
9.4884 | -89.8402
9.4887 | -89.8404
9.489 | -89.8405
9.4894 | -89.8416
9.4897 | -89.8424
9.49 | -89.8425
9.4904 | -89.8432
9.4907 | -89.8428
9.491 | -89.8429
9.4914 | -89.8434
9.4917 | -89.8449
9.492 | -89.8456
9.4923 | -89.8459
9.4927 | -89.8461
9.493 | -89.8463
9.4933 | -89.8467
9.4937 | -89.8471
9.494 | -89.8477
9.4943 | -89.848
9.4946 | -89.8479
9.495 | -89.8483
9.4953 | -89.8488
9.4956 | -89.8489
9.496 | -89.8494
9.4963 | -89.8491
9.4966 | -89.8498
9.4969 | -89.8506
9.4973 | -89.8509
9.4976 | -89.8514
9.4979 | -89.8515
9.4983 | -89.8516
9.4986 | -89.8518
9.4989 | -89.8521
9.4992 | -89.8522
9.4996 | -89.8524
9.4999 | -89.8527
9.5002 | -89.853
9.5006 | -89.8532
9.5009 | -89.8537
9.5012 | -89.8542
9.5016 | -89.8545
9.5019 | -89.8548
9.5022 | -89.8551
9.5025 | -89.8555
9.5029 | -89.8559
9.5032 | -89.8561
9.5035 | -89.8563
9.5039 | -89.857
9.5042 | -89.8574
9.5045 | -89.8577
9.5048 | -89.8583
9.5052 | -89.8591
9.5055 | -89.8596
9.5058 | -89.8602
9.5062 | -89.86
9.5065 | -89.86
9.5068 | -89.8604
9.5071 | -89.8605
9.5075 | -89.8607
9.5078 | -89.8619
9.5081 | -89.8624
9.5085 | -89.8629
9.5088 | -89.8634
9.5091 | -89.8637
9.5095 | -89.8637
9.5098 | -89.8642
9.5101 | -89.8647
9.5104 | -89.865
9.5108 | -89.865
9.5111 | -89.865
9.5114 | -89.8652
9.5118 | -89.8657
9.5121 | -89.866
9.5124 | -89.8665
9.5127 | -89.8668
9.5131 | -89.8669
9.5134 | -89.8679
9.5137 | -89.869
9.5141 | -89.8692
9.5144 | -89.8696
9.5147 | -89.8697
9.515 | -89.8702
9.5154 | -89.8703
9.5157 | -89.8711
9.516 | -89.8712
9.5164 | -89.8717
9.5167 | -89.8721
9.517 | -89.8724
9.5174 | -89.8728
9.5177 | -89.8736
9.518 | -89.8743
9.5183 | -89.8747
9.5187 | -89.8749
9.519 | -89.8751
9.5193 | -89.8755
9.5197 | -89.8759
9.52 | -89.8765
9.5203 | -89.8771
9.5206 | -89.8777
9.521 | -89.8785
9.5213 | -89.8789
9.5216 | -89.8804
9.522 | -89.8813
9.5223 | -89.8817
9.5226 | -89.8819
9.5229 | -89.8824
9.5233 | -89.8826
9.5236 | -89.8828
9.5239 | -89.8834
9.5243 | -89.8839
9.5246 | -89.8843
9.5249 | -89.8848
9.5253 | -89.8851
9.5256 | -89.8856
9.5259 | -89.8863
9.5262 | -89.8878
9.5266 | -89.8889
9.5269 | -89.8893
9.5272 | -89.8894
9.5276 | -89.8898
9.5279 | -89.8902
9.5282 | -89.8906
9.5285 | -89.8911
9.5289 | -89.8917
9.5292 | -89.8923
9.5295 | -89.8927
9.5299 | -89.8934
9.5302 | -89.894
9.5305 | -89.8944
9.5308 | -89.8949
9.5312 | -89.895
9.5315 | -89.8952
9.5318 | -89.8955
9.5322 | -89.896
9.5325 | -89.8962
9.5328 | -89.8965
9.5332 | -89.897
9.5335 | -89.8975
9.5338 | -89.8979
9.5341 | -89.8986
9.5345 | -89.8992
9.5348 | -89.8997
9.5351 | -89.9003
9.5355 | -89.901
9.5358 | -89.9023
9.5361 | -89.9028
9.5364 | -89.9033
9.5368 | -89.9038
9.5371 | -89.9046
9.5374 | -89.9053
9.5378 | -89.9058
9.5381 | -89.906
9.5384 | -89.9058
9.5387 | -89.9055
9.5391 | -89.9058
9.5394 | -89.9067
9.5397 | -89.9072
9.5401 | -89.908
9.5404 | -89.9089
9.5407 | -89.9094
9.5411 | -89.91
9.5414 | -89.9108
9.5417 | -89.9114
9.542 | -89.9121
9.5424 | -89.9127
9.5427 | -89.9131
9.543 | -89.9145
9.5434 | -89.916
9.5437 | -89.9157
9.544 | -89.9156
9.5443 | -89.9164
9.5447 | -89.917
9.545 | -89.9176
9.5453 | -89.918
9.5457 | -89.9187
9.546 | -89.92
9.5463 | -89.9199
9.5466 | -89.9193
9.547 | -89.9193
9.5473 | -89.9198
9.5476 | -89.9202
9.548 | -89.9217
9.5483 | -89.9226
9.5486 | -89.923
9.5489 | -89.9238
9.5493 | -89.9243
9.5496 | -89.925
9.5499 | -89.926
9.5503 | -89.9263
9.5506 | -89.9267
9.5509 | -89.9271
9.5513 | -89.9275
9.5516 | -89.9282
9.5519 | -89.9288
9.5522 | -89.9294
9.5526 | -89.93
9.5529 | -89.9305
9.5532 | -89.9312
9.5536 | -89.9321
9.5539 | -89.9326
9.5542 | -89.9335
9.5545 | -89.9344
9.5549 | -89.9349
9.5552 | -89.9354
9.5555 | -89.9359
9.5559 | -89.9364
9.5562 | -89.937
9.5565 | -89.9374
9.5568 | -89.9379
9.5572 | -89.9385
9.5575 | -89.9392
9.5578 | -89.9403
9.5582 | -89.9411
9.5585 | -89.9416
9.5588 | -89.9421
9.5592 | -89.9426
9.5595 | -89.9431
9.5598 | -89.9432
9.5601 | -89.9438
9.5605 | -89.9442
9.5608 | -89.9445
9.5611 | -89.9456
9.5615 | -89.9463
9.5618 | -89.9468
9.5621 | -89.9474
9.5624 | -89.9482
9.5628 | -89.9489
9.5631 | -89.9494
9.5634 | -89.9501
9.5638 | -89.9506
9.5641 | -89.9511
9.5644 | -89.9514
9.5647 | -89.9514
9.5651 | -89.9518
9.5654 | -89.9521
9.5657 | -89.9527
9.5661 | -89.9527
9.5664 | -89.952
9.5667 | -89.9523
9.5671 | -89.9529
9.5674 | -89.9536
9.5677 | -89.954
9.568 | -89.9544
9.5684 | -89.9549
9.5687 | -89.9554
9.569 | -89.9558
9.5694 | -89.956
9.5697 | -89.9564
9.57 | -89.9569
9.5703 | -89.9571
9.5707 | -89.9576
9.571 | -89.9576
9.5713 | -89.9575
9.5717 | -89.9581
9.572 | -89.9585
9.5723 | -89.9592
9.5726 | -89.9597
9.573 | -89.9602
9.5733 | -89.9611
9.5736 | -89.9615
9.574 | -89.962
9.5743 | -89.9626
9.5746 | -89.9634
9.575 | -89.9643
9.5753 | -89.9649
9.5756 | -89.9656
9.5759 | -89.9663
9.5763 | -89.9665
9.5766 | -89.9665
9.5769 | -89.9676
9.5773 | -89.9679
9.5776 | -89.9679
9.5779 | -89.9682
9.5782 | -89.9686
9.5786 | -89.9693
9.5789 | -89.97
9.5792 | -89.9708
9.5796 | -89.9714
9.5799 | -89.9723
9.5802 | -89.9725
9.5805 | -89.9727
9.5809 | -89.9735
9.5812 | -89.9742
9.5815 | -89.974
9.5819 | -89.9738
9.5822 | -89.9744
9.5825 | -89.9749
9.5829 | -89.9756
9.5832 | -89.9761
9.5835 | -89.9765
9.5838 | -89.9776
9.5842 | -89.9785
9.5845 | -89.9791
9.5848 | -89.9797
9.5852 | -89.9795
9.5855 | -89.9795
9.5858 | -89.9801
9.5861 | -89.9817
9.5865 | -89.9843
9.5868 | -89.9856
9.5871 | -89.9862
9.5875 | -89.9868
9.5878 | -89.9872
9.5881 | -89.9879
9.5884 | -89.9887
9.5888 | -89.9891
9.5891 | -89.9895
9.5894 | -89.9901
9.5898 | -89.991
9.5901 | -89.992
9.5904 | -89.9935
9.5907 | -89.9942
9.5911 | -89.9947
9.5914 | -89.9952
9.5917 | -89.9962
9.5921 | -89.9965
9.5924 | -89.997
9.5927 | -89.9974
9.5931 | -89.9982
9.5934 | -89.9988
9.5937 | -89.9996
9.594 | -90.0001
9.5944 | -90.0005
9.5947 | -90.0011
9.595 | -90.0018
9.5954 | -90.0026
9.5957 | -90.003
9.596 | -90.0035
9.5963 | -90.0042
9.5967 | -90.0053
9.597 | -90.0061
9.5973 | -90.006
9.5977 | -90.0066
9.598 | -90.0072
9.5983 | -90.0079
9.5986 | -90.0081
9.599 | -90.0085
9.5993 | -90.0091
9.5996 | -90.0102
9.6 | -90.011
9.6003 | -90.0112
9.6006 | -90.0118
9.601 | -90.0117
9.6013 | -90.0129
9.6016 | -90.0135
9.6019 | -90.0141
9.6023 | -90.015
9.6026 | -90.0156
9.6029 | -90.0164
9.6033 | -90.0168
9.6036 | -90.0167
9.6039 | -90.0173
9.6042 | -90.0176
9.6046 | -90.0182
9.6049 | -90.0186
9.6052 | -90.0189
9.6056 | -90.0194
9.6059 | -90.0199
9.6062 | -90.0205
9.6065 | -90.0208
9.6069 | -90.0214
9.6072 | -90.022
9.6075 | -90.0225
9.6079 | -90.0236
9.6082 | -90.0244
9.6085 | -90.0249
9.6089 | -90.0254
9.6092 | -90.0261
9.6095 | -90.0267
9.6098 | -90.0269
9.6102 | -90.0274
9.6105 | -90.0279
9.6108 | -90.0281
9.6112 | -90.0289
9.6115 | -90.0296
9.6118 | -90.0301
9.6121 | -90.0305
9.6125 | -90.0312
9.6128 | -90.0312
9.6131 | -90.0314
9.6135 | -90.032
9.6138 | -90.0328
9.6141 | -90.0337
9.6144 | -90.0345
9.6148 | -90.0353
9.6151 | -90.0358
9.6154 | -90.0359
9.6158 | -90.0365
9.6161 | -90.0371
9.6164 | -90.0376
9.6168 | -90.0378
9.6171 | -90.0384
9.6174 | -90.0389
9.6177 | -90.0392
9.6181 | -90.0395
9.6184 | -90.0401
9.6187 | -90.0406
9.6191 | -90.0395
9.6194 | -90.0385
9.6197 | -90.039
9.62 | -90.0395
9.6204 | -90.0404
9.6207 | -90.0413
9.621 | -90.0421
9.6214 | -90.0426
9.6217 | -90.043
9.622 | -90.0428
9.6223 | -90.0422
9.6227 | -90.0418
9.623 | -90.0424
9.6233 | -90.0427
9.6237 | -90.0433
9.624 | -90.0433
9.6243 | -90.0434
9.6247 | -90.0443
9.625 | -90.045
9.6253 | -90.0454
9.6256 | -90.0458
9.626 | -90.0463
9.6263 | -90.0468
9.6266 | -90.0473
9.627 | -90.0481
9.6273 | -90.049
9.6276 | -90.0493
9.6279 | -90.0496
9.6283 | -90.0503
9.6286 | -90.0509
9.6289 | -90.0512
9.6293 | -90.0519
9.6296 | -90.0526
9.6299 | -90.0527
9.6302 | -90.0532
9.6306 | -90.0554
9.6309 | -90.0572
9.6312 | -90.0581
9.6316 | -90.0582
9.6319 | -90.0586
9.6322 | -90.0591
9.6326 | -90.0598
9.6329 | -90.0608
9.6332 | -90.0613
9.6335 | -90.0614
9.6339 | -90.0618
9.6342 | -90.0626
9.6345 | -90.0633
9.6349 | -90.0637
9.6352 | -90.0645
9.6355 | -90.0646
9.6358 | -90.0652
9.6362 | -90.0662
9.6365 | -90.0666
9.6368 | -90.0671
9.6372 | -90.0671
9.6375 | -90.0675
9.6378 | -90.0681
9.6381 | -90.0687
9.6385 | -90.0694
9.6388 | -90.07
9.6391 | -90.0709
9.6395 | -90.0712
9.6398 | -90.0716
9.6401 | -90.0725
9.6404 | -90.0733
9.6408 | -90.0736
9.6411 | -90.0741
9.6414 | -90.0744
9.6418 | -90.075
9.6421 | -90.0755
9.6424 | -90.0758
9.6428 | -90.0765
9.6431 | -90.0771
9.6434 | -90.0773
9.6437 | -90.0775
9.6441 | -90.0784
9.6444 | -90.0793
9.6447 | -90.0798
9.6451 | -90.0801
9.6454 | -90.0809
9.6457 | -90.0815
9.646 | -90.0818
9.6464 | -90.0824
9.6467 | -90.0829
9.647 | -90.083
9.6474 | -90.0831
9.6477 | -90.0836
9.648 | -90.0847
9.6483 | -90.0851
9.6487 | -90.0862
9.649 | -90.0861
9.6493 | -90.0863
9.6497 | -90.087
9.65 | -90.0881
9.6503 | -90.089
9.6507 | -90.0897
9.651 | -90.0903
9.6513 | -90.0909
9.6516 | -90.0913
9.652 | -90.0916
9.6523 | -90.092
9.6526 | -90.092
9.653 | -90.0923
9.6533 | -90.0932
9.6536 | -90.0947
9.6539 | -90.0961
9.6543 | -90.0969
9.6546 | -90.0973
9.6549 | -90.0979
9.6553 | -90.0982
9.6556 | -90.098
9.6559 | -90.0982
9.6562 | -90.0989
9.6566 | -90.1001
9.6569 | -90.1009
9.6572 | -90.1014
9.6576 | -90.1017
9.6579 | -90.1027
9.6582 | -90.1047
9.6586 | -90.1054
9.6589 | -90.1059
9.6592 | -90.1072
9.6595 | -90.1086
9.6599 | -90.1093
9.6602 | -90.11
9.6605 | -90.1103
9.6609 | -90.1108
9.6612 | -90.1111
9.6615 | -90.1107
9.6618 | -90.1109
9.6622 | -90.1111
9.6625 | -90.1118
9.6628 | -90.1125
9.6632 | -90.1133
9.6635 | -90.1141
9.6638 | -90.1149
9.6641 | -90.1152
9.6645 | -90.1154
9.6648 | -90.1156
9.6651 | -90.1163
9.6655 | -90.1168
9.6658 | -90.1177
9.6661 | -90.1185
9.6665 | -90.1191
9.6668 | -90.1202
9.6671 | -90.1211
9.6674 | -90.1216
9.6678 | -90.1216
9.6681 | -90.1221
9.6684 | -90.1228
9.6688 | -90.1234
9.6691 | -90.1242
9.6694 | -90.1248
9.6697 | -90.1253
9.6701 | -90.1256
9.6704 | -90.126
9.6707 | -90.1265
9.6711 | -90.1269
9.6714 | -90.1273
9.6717 | -90.1273
9.672 | -90.1275
9.6724 | -90.1278
9.6727 | -90.1281
9.673 | -90.1289
9.6734 | -90.13
9.6737 | -90.1306
9.674 | -90.1311
9.6744 | -90.1314
9.6747 | -90.1321
9.675 | -90.1326
9.6753 | -90.1332
9.6757 | -90.1338
9.676 | -90.1343
9.6763 | -90.1348
9.6767 | -90.1352
9.677 | -90.136
9.6773 | -90.1368
9.6776 | -90.1379
9.678 | -90.1391
9.6783 | -90.14
9.6786 | -90.1405
9.679 | -90.141
9.6793 | -90.1419
9.6796 | -90.1424
9.6799 | -90.143
9.6803 | -90.1436
9.6806 | -90.1449
9.6809 | -90.1455
9.6813 | -90.1457
9.6816 | -90.146
9.6819 | -90.1464
9.6822 | -90.1468
9.6826 | -90.148
9.6829 | -90.1503
9.6832 | -90.152
9.6836 | -90.1532
9.6839 | -90.1537
9.6842 | -90.1529
9.6846 | -90.1524
9.6849 | -90.1529
9.6852 | -90.1533
9.6855 | -90.1537
9.6859 | -90.1543
9.6862 | -90.1547
9.6865 | -90.1554
9.6869 | -90.1557
9.6872 | -90.1559
9.6875 | -90.1561
9.6878 | -90.1565
9.6882 | -90.1572
9.6885 | -90.1579
9.6888 | -90.1585
9.6892 | -90.1588
9.6895 | -90.1594
9.6898 | -90.1597
9.6901 | -90.1602
9.6905 | -90.1607
9.6908 | -90.1613
9.6911 | -90.162
9.6915 | -90.1629
9.6918 | -90.1637
9.6921 | -90.1642
9.6925 | -90.1647
9.6928 | -90.1651
9.6931 | -90.1659
9.6934 | -90.1674
9.6938 | -90.1677
9.6941 | -90.1678
9.6944 | -90.1683
9.6948 | -90.169
9.6951 | -90.1699
9.6954 | -90.1706
9.6957 | -90.1703
9.6961 | -90.1704
9.6964 | -90.1705
9.6967 | -90.1704
9.6971 | -90.171
9.6974 | -90.1716
9.6977 | -90.1724
9.698 | -90.1732
9.6984 | -90.174
9.6987 | -90.1747
9.699 | -90.1746
9.6994 | -90.1739
9.6997 | -90.174
9.7 | -90.1751
9.7004 | -90.1757
9.7007 | -90.1765
9.701 | -90.1771
9.7013 | -90.1777
9.7017 | -90.1783
9.702 | -90.1787
9.7023 | -90.1792
9.7027 | -90.179
9.703 | -90.1789
9.7033 | -90.1794
9.7036 | -90.1799
9.704 | -90.1801
9.7043 | -90.181
9.7046 | -90.1828
9.705 | -90.1854
9.7053 | -90.1858
9.7056 | -90.1866
9.7059 | -90.1871
9.7063 | -90.1874
9.7066 | -90.1881
9.7069 | -90.1884
9.7073 | -90.1883
9.7076 | -90.1888
9.7079 | -90.1894
9.7083 | -90.1899
9.7086 | -90.1907
9.7089 | -90.1915
9.7092 | -90.1919
9.7096 | -90.1924
9.7099 | -90.1928
9.7102 | -90.1932
9.7106 | -90.1933
9.7109 | -90.1938
9.7112 | -90.1942
9.7115 | -90.1944
9.7119 | -90.1945
9.7122 | -90.1948
9.7125 | -90.1952
9.7129 | -90.1965
9.7132 | -90.1973
9.7135 | -90.1974
9.7138 | -90.1978
9.7142 | -90.1983
9.7145 | -90.1987
9.7148 | -90.1994
9.7152 | -90.1999
9.7155 | -90.2002
9.7158 | -90.2003
9.7162 | -90.201
9.7165 | -90.2011
9.7168 | -90.2014
9.7171 | -90.2019
9.7175 | -90.2025
9.7178 | -90.203
9.7181 | -90.2035
9.7185 | -90.204
9.7188 | -90.2044
9.7191 | -90.2044
9.7194 | -90.2053
9.7198 | -90.2064
9.7201 | -90.2064
9.7204 | -90.2065
9.7208 | -90.2071
9.7211 | -90.2076
9.7214 | -90.2083
9.7217 | -90.2091
9.7221 | -90.21
9.7224 | -90.2108
9.7227 | -90.2113
9.7231 | -90.2119
9.7234 | -90.2121
9.7237 | -90.2125
9.7241 | -90.213
9.7244 | -90.2137
9.7247 | -90.2145
9.725 | -90.2149
9.7254 | -90.2153
9.7257 | -90.2157
9.726 | -90.2162
9.7264 | -90.2167
9.7267 | -90.2173
9.727 | -90.2179
9.7273 | -90.2184
9.7277 | -90.2188
9.728 | -90.2193
9.7283 | -90.2196
9.7287 | -90.2202
9.729 | -90.2208
9.7293 | -90.2211
9.7296 | -90.2219
9.73 | -90.2226
9.7303 | -90.2229
9.7306 | -90.2232
9.731 | -90.2236
9.7313 | -90.2241
9.7316 | -90.2241
9.7319 | -90.2244
9.7323 | -90.2249
9.7326 | -90.2254
9.7329 | -90.2263
9.7333 | -90.2269
9.7336 | -90.2277
9.7339 | -90.2282
9.7343 | -90.2283
9.7346 | -90.229
9.7349 | -90.2299
9.7352 | -90.2305
9.7356 | -90.2311
9.7359 | -90.2312
9.7362 | -90.2314
9.7366 | -90.2321
9.7369 | -90.2328
9.7372 | -90.2335
9.7375 | -90.2342
9.7379 | -90.2345
9.7382 | -90.2349
9.7385 | -90.2354
9.7389 | -90.2354
9.7392 | -90.2359
9.7395 | -90.2363
9.7398 | -90.237
9.7402 | -90.2379
9.7405 | -90.2383
9.7408 | -90.2389
9.7412 | -90.2399
9.7415 | -90.2408
9.7418 | -90.2412
9.7422 | -90.2417
9.7425 | -90.2418
9.7428 | -90.2418
9.7431 | -90.2421
9.7435 | -90.2424
9.7438 | -90.2429
9.7441 | -90.2431
9.7445 | -90.2436
9.7448 | -90.2445
9.7451 | -90.2451
9.7454 | -90.2451
9.7458 | -90.2456
9.7461 | -90.246
9.7464 | -90.2465
9.7468 | -90.2472
9.7471 | -90.2479
9.7474 | -90.2485
9.7477 | -90.2487
9.7481 | -90.2496
9.7484 | -90.2501
9.7487 | -90.2503
9.7491 | -90.251
9.7494 | -90.252
9.7497 | -90.2525
9.7501 | -90.2527
9.7504 | -90.2534
9.7507 | -90.2546
9.751 | -90.2551
9.7514 | -90.255
9.7517 | -90.2556
9.752 | -90.2564
9.7524 | -90.2568
9.7527 | -90.2572
9.753 | -90.2577
9.7533 | -90.2583
9.7537 | -90.2597
9.754 | -90.2597
9.7543 | -90.2593
9.7547 | -90.2595
9.755 | -90.2603
9.7553 | -90.2605
9.7556 | -90.2605
9.756 | -90.2605
9.7563 | -90.2613
9.7566 | -90.2619
9.757 | -90.2624
9.7573 | -90.2627
9.7576 | -90.263
9.758 | -90.2634
9.7583 | -90.2637
9.7586 | -90.2645
9.7589 | -90.2652
9.7593 | -90.266
9.7596 | -90.2676
9.7599 | -90.269
9.7603 | -90.2693
9.7606 | -90.2696
9.7609 | -90.2709
9.7612 | -90.2717
9.7616 | -90.2721
9.7619 | -90.2723
9.7622 | -90.2725
9.7626 | -90.2728
9.7629 | -90.2728
9.7632 | -90.2732
9.7635 | -90.2735
9.7639 | -90.2737
9.7642 | -90.275
9.7645 | -90.2766
9.7649 | -90.2776
9.7652 | -90.278
9.7655 | -90.2787
9.7659 | -90.2795
9.7662 | -90.2798
9.7665 | -90.2798
9.7668 | -90.2803
9.7672 | -90.2808
9.7675 | -90.2816
9.7678 | -90.2823
9.7682 | -90.2825
9.7685 | -90.2828
9.7688 | -90.2834
9.7691 | -90.2837
9.7695 | -90.2841
9.7698 | -90.2848
9.7701 | -90.2855
9.7705 | -90.2862
9.7708 | -90.2866
9.7711 | -90.2873
9.7714 | -90.2877
9.7718 | -90.2884
9.7721 | -90.289
9.7724 | -90.2895
9.7728 | -90.2899
9.7731 | -90.2909
9.7734 | -90.2919
9.7738 | -90.2925
9.7741 | -90.2931
9.7744 | -90.2934
9.7747 | -90.294
9.7751 | -90.2944
9.7754 | -90.2943
9.7757 | -90.2943
9.7761 | -90.2947
9.7764 | -90.2949
9.7767 | -90.295
9.777 | -90.2956
9.7774 | -90.296
9.7777 | -90.2961
9.778 | -90.2966
9.7784 | -90.2969
9.7787 | -90.2972
9.779 | -90.2976
9.7793 | -90.2977
9.7797 | -90.2982
9.78 | -90.2987
9.7803 | -90.2987
9.7807 | -90.299
9.781 | -90.2994
9.7813 | -90.3
9.7817 | -90.3002
9.782 | -90.3007
9.7823 | -90.3011
9.7826 | -90.3013
9.783 | -90.3018
9.7833 | -90.3021
9.7836 | -90.3021
9.784 | -90.3023
9.7843 | -90.3027
9.7846 | -90.3032
9.7849 | -90.3034
9.7853 | -90.3037
9.7856 | -90.3041
9.7859 | -90.3043
9.7863 | -90.3045
9.7866 | -90.3048
9.7869 | -90.3054
9.7872 | -90.3056
9.7876 | -90.3056
9.7879 | -90.3055
9.7882 | -90.306
9.7886 | -90.3066
9.7889 | -90.3071
9.7892 | -90.3076
9.7895 | -90.3087
9.7899 | -90.3091
9.7902 | -90.3097
9.7905 | -90.3102
9.7909 | -90.3111
9.7912 | -90.3115
9.7915 | -90.3124
9.7919 | -90.3136
9.7922 | -90.3142
9.7925 | -90.3151
9.7928 | -90.317
9.7932 | -90.318
9.7935 | -90.3183
9.7938 | -90.3187
9.7942 | -90.319
9.7945 | -90.3198
9.7948 | -90.3201
9.7951 | -90.3206
9.7955 | -90.3215
9.7958 | -90.3222
9.7961 | -90.323
9.7965 | -90.3245
9.7968 | -90.3253
9.7971 | -90.3257
9.7974 | -90.326
9.7978 | -90.3256
9.7981 | -90.3257
9.7984 | -90.3259
9.7988 | -90.3262
9.7991 | -90.326
9.7994 | -90.3253
9.7998 | -90.3254
9.8001 | -90.3252
9.8004 | -90.3251
9.8007 | -90.3248
9.8011 | -90.3254
9.8014 | -90.3261
9.8017 | -90.3264
9.8021 | -90.3266
9.8024 | -90.3268
9.8027 | -90.3269
9.803 | -90.3271
9.8034 | -90.3274
9.8037 | -90.328
9.804 | -90.3286
9.8044 | -90.3289
9.8047 | -90.3292
9.805 | -90.3296
9.8053 | -90.3298
9.8057 | -90.3302
9.806 | -90.3304
9.8063 | -90.331
9.8067 | -90.3318
9.807 | -90.3316
9.8073 | -90.3314
9.8077 | -90.3318
9.808 | -90.3323
9.8083 | -90.3329
9.8086 | -90.333
9.809 | -90.3332
9.8093 | -90.3335
9.8096 | -90.3332
9.81 | -90.3334
9.8103 | -90.3336
9.8106 | -90.3338
9.8109 | -90.3342
9.8113 | -90.3349
9.8116 | -90.3359
9.8119 | -90.3366
9.8123 | -90.3374
9.8126 | -90.3383
9.8129 | -90.339
9.8132 | -90.3392
9.8136 | -90.3394
9.8139 | -90.3398
9.8142 | -90.34
9.8146 | -90.3409
9.8149 | -90.3427
9.8152 | -90.3442
9.8156 | -90.3445
9.8159 | -90.345
9.8162 | -90.3456
9.8165 | -90.3463
9.8169 | -90.3466
9.8172 | -90.347
9.8175 | -90.3473
9.8179 | -90.3475
9.8182 | -90.3477
9.8185 | -90.3484
9.8188 | -90.3492
9.8192 | -90.3496
9.8195 | -90.3502
9.8198 | -90.3502
9.8202 | -90.3503
9.8205 | -90.3505
9.8208 | -90.3506
9.8211 | -90.3503
9.8215 | -90.3503
9.8218 | -90.3506
9.8221 | -90.3508
9.8225 | -90.3511
9.8228 | -90.3513
9.8231 | -90.3517
9.8234 | -90.3521
9.8238 | -90.3517
9.8241 | -90.3523
9.8244 | -90.353
9.8248 | -90.3537
9.8251 | -90.354
9.8254 | -90.3535
9.8258 | -90.3539
9.8261 | -90.3542
9.8264 | -90.3545
9.8267 | -90.3549
9.8271 | -90.3553
9.8274 | -90.3555
9.8277 | -90.3555
9.8281 | -90.3557
9.8284 | -90.3559
9.8287 | -90.3563
9.829 | -90.357
9.8294 | -90.3576
9.8297 | -90.3577
9.83 | -90.358
9.8304 | -90.3583
9.8307 | -90.3586
9.831 | -90.3591
9.8314 | -90.3593
9.8317 | -90.3602
9.832 | -90.3605
9.8323 | -90.3609
9.8327 | -90.3612
9.833 | -90.3613
9.8333 | -90.3615
9.8337 | -90.3616
9.834 | -90.3614
9.8343 | -90.3615
9.8346 | -90.3618
9.835 | -90.3619
9.8353 | -90.3627
9.8356 | -90.3632
9.836 | -90.3635
9.8363 | -90.3641
9.8366 | -90.3642
9.8369 | -90.3647
9.8373 | -90.3652
9.8376 | -90.3655
9.8379 | -90.3656
9.8383 | -90.3651
9.8386 | -90.3655
9.8389 | -90.3656
9.8392 | -90.366
9.8396 | -90.3663
9.8399 | -90.3664
9.8402 | -90.3666
9.8406 | -90.3669
9.8409 | -90.3674
9.8412 | -90.3682
9.8416 | -90.3687
9.8419 | -90.3699
9.8422 | -90.3707
9.8425 | -90.3709
9.8429 | -90.3712
9.8432 | -90.3716
9.8435 | -90.3721
9.8439 | -90.3726
9.8442 | -90.373
9.8445 | -90.3731
9.8448 | -90.3733
9.8452 | -90.3734
9.8455 | -90.3735
9.8458 | -90.3737
9.8462 | -90.3741
9.8465 | -90.3754
9.8468 | -90.3762
9.8471 | -90.3766
9.8475 | -90.3755
9.8478 | -90.3754
9.8481 | -90.3758
9.8485 | -90.3761
9.8488 | -90.3764
9.8491 | -90.3768
9.8495 | -90.3777
9.8498 | -90.3783
9.8501 | -90.3787
9.8504 | -90.3791
9.8508 | -90.3791
9.8511 | -90.3791
9.8514 | -90.3792
9.8518 | -90.3796
9.8521 | -90.3805
9.8524 | -90.3816
9.8527 | -90.3822
9.8531 | -90.3825
9.8534 | -90.3824
9.8537 | -90.3827
9.8541 | -90.3834
9.8544 | -90.3837
9.8547 | -90.3842
9.855 | -90.3845
9.8554 | -90.3847
9.8557 | -90.3852
9.856 | -90.3856
9.8564 | -90.3854
9.8567 | -90.3853
9.857 | -90.3855
9.8574 | -90.386
9.8577 | -90.3858
9.858 | -90.3855
9.8583 | -90.3853
9.8587 | -90.3853
9.859 | -90.3855
9.8593 | -90.3857
9.8597 | -90.3861
9.86 | -90.3863
9.8603 | -90.3867
9.8606 | -90.3873
9.861 | -90.3878
9.8613 | -90.3881
9.8616 | -90.3884
9.862 | -90.3883
9.8623 | -90.3883
9.8626 | -90.3886
9.8629 | -90.3888
9.8633 | -90.389
9.8636 | -90.3895
9.8639 | -90.39
9.8643 | -90.3905
9.8646 | -90.3909
9.8649 | -90.391
9.8653 | -90.3909
9.8656 | -90.3912
9.8659 | -90.3912
9.8662 | -90.3915
9.8666 | -90.3914
9.8669 | -90.3912
9.8672 | -90.3916
9.8676 | -90.3923
9.8679 | -90.3931
9.8682 | -90.3943
9.8685 | -90.3949
9.8689 | -90.3949
9.8692 | -90.3951
9.8695 | -90.3956
9.8699 | -90.3959
9.8702 | -90.3961
9.8705 | -90.396
9.8708 | -90.3958
9.8712 | -90.396
9.8715 | -90.3961
9.8718 | -90.3962
9.8722 | -90.3974
9.8725 | -90.3981
9.8728 | -90.3985
9.8732 | -90.3989
9.8735 | -90.3994
9.8738 | -90.3998
9.8741 | -90.4001
9.8745 | -90.4003
9.8748 | -90.3999
9.8751 | -90.4
9.8755 | -90.4003
9.8758 | -90.4006
9.8761 | -90.4007
9.8764 | -90.4013
9.8768 | -90.4016
9.8771 | -90.4019
9.8774 | -90.4016
9.8778 | -90.4015
9.8781 | -90.4022
9.8784 | -90.4029
9.8787 | -90.4028
9.8791 | -90.4027
9.8794 | -90.4029
9.8797 | -90.4041
9.8801 | -90.4061
9.8804 | -90.4067
9.8807 | -90.4073
9.8811 | -90.408
9.8814 | -90.4086
9.8817 | -90.4091
9.882 | -90.4092
9.8824 | -90.4093
9.8827 | -90.4103
9.883 | -90.4111
9.8834 | -90.4113
9.8837 | -90.4119
9.884 | -90.4122
9.8843 | -90.4129
9.8847 | -90.4133
9.885 | -90.4136
9.8853 | -90.4133
9.8857 | -90.413
9.886 | -90.4125
9.8863 | -90.4127
9.8866 | -90.4132
9.887 | -90.4137
9.8873 | -90.4135
9.8876 | -90.4137
9.888 | -90.4138
9.8883 | -90.414
9.8886 | -90.4145
9.8889 | -90.4146
9.8893 | -90.4141
9.8896 | -90.4143
9.8899 | -90.4144
9.8903 | -90.4149
9.8906 | -90.415
9.8909 | -90.415
9.8913 | -90.4151
9.8916 | -90.4149
9.8919 | -90.4145
9.8922 | -90.4143
9.8926 | -90.4148
9.8929 | -90.4157
9.8932 | -90.416
9.8936 | -90.4161
9.8939 | -90.4163
9.8942 | -90.4164
9.8945 | -90.4159
9.8949 | -90.4153
9.8952 | -90.4153
9.8955 | -90.4157
9.8959 | -90.4152
9.8962 | -90.4145
9.8965 | -90.4144
9.8968 | -90.4149
9.8972 | -90.4155
9.8975 | -90.4158
9.8978 | -90.416
9.8982 | -90.4166
9.8985 | -90.4167
9.8988 | -90.417
9.8992 | -90.4174
9.8995 | -90.4174
9.8998 | -90.4172
9.9001 | -90.4173
9.9005 | -90.4175
9.9008 | -90.4181
9.9011 | -90.4187
9.9015 | -90.4191
9.9018 | -90.4193
9.9021 | -90.4192
9.9024 | -90.4192
9.9028 | -90.4194
9.9031 | -90.4197
9.9034 | -90.4199
9.9038 | -90.42
9.9041 | -90.4203
9.9044 | -90.4204
9.9047 | -90.4205
9.9051 | -90.4206
9.9054 | -90.4206
9.9057 | -90.4209
9.9061 | -90.4218
9.9064 | -90.422
9.9067 | -90.4226
9.9071 | -90.4227
9.9074 | -90.4229
9.9077 | -90.4229
9.908 | -90.423
9.9084 | -90.4232
9.9087 | -90.4239
9.909 | -90.4245
9.9094 | -90.4252
9.9097 | -90.4256
9.91 | -90.4254
9.9103 | -90.4254
9.9107 | -90.4257
9.911 | -90.4266
9.9113 | -90.4274
9.9117 | -90.4278
9.912 | -90.4285
9.9123 | -90.4288
9.9126 | -90.4289
9.913 | -90.4291
9.9133 | -90.4292
9.9136 | -90.4294
9.914 | -90.4301
9.9143 | -90.4302
9.9146 | -90.4303
9.915 | -90.43
9.9153 | -90.4302
9.9156 | -90.4305
9.9159 | -90.4306
9.9163 | -90.4305
9.9166 | -90.4303
9.9169 | -90.4308
9.9173 | -90.4315
9.9176 | -90.4315
9.9179 | -90.432
9.9182 | -90.4323
9.9186 | -90.4322
9.9189 | -90.4322
9.9192 | -90.4319
9.9196 | -90.4322
9.9199 | -90.4324
9.9202 | -90.4327
9.9205 | -90.4325
9.9209 | -90.4318
9.9212 | -90.4314
9.9215 | -90.4318
9.9219 | -90.4319
9.9222 | -90.4321
9.9225 | -90.4326
9.9229 | -90.4332
9.9232 | -90.4332
9.9235 | -90.4337
9.9238 | -90.4341
9.9242 | -90.4349
9.9245 | -90.4354
9.9248 | -90.4355
9.9252 | -90.4354
9.9255 | -90.4358
9.9258 | -90.4364
9.9261 | -90.4367
9.9265 | -90.4371
9.9268 | -90.4373
9.9271 | -90.4376
9.9275 | -90.4378
9.9278 | -90.4387
9.9281 | -90.4395
9.9284 | -90.4401
9.9288 | -90.4406
9.9291 | -90.4406
9.9294 | -90.4408
9.9298 | -90.4411
9.9301 | -90.4413
9.9304 | -90.4418
9.9307 | -90.4423
9.9311 | -90.4424
9.9314 | -90.4427
9.9317 | -90.4435
9.9321 | -90.4436
9.9324 | -90.4435
9.9327 | -90.4433
9.9331 | -90.4437
9.9334 | -90.4439
9.9337 | -90.4443
9.934 | -90.4444
9.9344 | -90.4448
9.9347 | -90.4452
9.935 | -90.4455
9.9354 | -90.4459
9.9357 | -90.4471
9.936 | -90.4481
9.9363 | -90.4478
9.9367 | -90.448
9.937 | -90.4482
9.9373 | -90.4482
9.9377 | -90.4485
9.938 | -90.4486
9.9383 | -90.4486
9.9386 | -90.4493
9.939 | -90.4499
9.9393 | -90.4502
9.9396 | -90.4504
9.94 | -90.4505
9.9403 | -90.4494
9.9406 | -90.448
9.941 | -90.4477
9.9413 | -90.448
9.9416 | -90.4481
9.9419 | -90.4483
9.9423 | -90.4492
9.9426 | -90.4498
9.9429 | -90.4502
9.9433 | -90.4502
9.9436 | -90.4501
9.9439 | -90.4503
9.9442 | -90.4505
9.9446 | -90.4512
9.9449 | -90.4515
9.9452 | -90.4515
9.9456 | -90.4516
9.9459 | -90.4517
9.9462 | -90.4516
9.9465 | -90.4518
9.9469 | -90.4525
9.9472 | -90.4524
9.9475 | -90.4525
9.9479 | -90.4527
9.9482 | -90.4528
9.9485 | -90.4535
9.9489 | -90.4535
9.9492 | -90.4533
9.9495 | -90.4531
9.9498 | -90.4531
9.9502 | -90.4534
9.9505 | -90.4536
9.9508 | -90.4542
9.9512 | -90.4541
9.9515 | -90.4542
9.9518 | -90.4547
9.9521 | -90.4554
9.9525 | -90.4557
9.9528 | -90.4556
9.9531 | -90.4558
9.9535 | -90.4561
9.9538 | -90.4566
9.9541 | -90.4568
9.9544 | -90.4572
9.9548 | -90.4577
9.9551 | -90.458
9.9554 | -90.4579
9.9558 | -90.4581
9.9561 | -90.4586
9.9564 | -90.4596
9.9568 | -90.4614
9.9571 | -90.4627
9.9574 | -90.4633
9.9577 | -90.4636
9.9581 | -90.4639
9.9584 | -90.4642
9.9587 | -90.4645
9.9591 | -90.4654
9.9594 | -90.4664
9.9597 | -90.4666
9.96 | -90.4668
9.9604 | -90.4671
9.9607 | -90.4673
9.961 | -90.4676
9.9614 | -90.4677
9.9617 | -90.4675
9.962 | -90.4684
9.9623 | -90.4685
9.9627 | -90.4687
9.963 | -90.4691
9.9633 | -90.4695
9.9637 | -90.4697
9.964 | -90.4695
9.9643 | -90.4698
9.9647 | -90.4705
9.965 | -90.4707
9.9653 | -90.4706
9.9656 | -90.4709
9.966 | -90.4712
9.9663 | -90.4716
9.9666 | -90.4723
9.967 | -90.4732
9.9673 | -90.4736
9.9676 | -90.4739
9.9679 | -90.4744
9.9683 | -90.4755
9.9686 | -90.4763
9.9689 | -90.4766
9.9693 | -90.4769
9.9696 | -90.4776
9.9699 | -90.4786
9.9702 | -90.4791
9.9706 | -90.4793
9.9709 | -90.4793
9.9712 | -90.48
9.9716 | -90.4802
9.9719 | -90.4805
9.9722 | -90.4808
9.9726 | -90.481
9.9729 | -90.4814
9.9732 | -90.4823
9.9735 | -90.4825
9.9739 | -90.4827
9.9742 | -90.4829
9.9745 | -90.4832
9.9749 | -90.4834
9.9752 | -90.484
9.9755 | -90.4851
9.9758 | -90.4861
9.9762 | -90.4866
9.9765 | -90.4866
9.9768 | -90.4868
9.9772 | -90.4872
9.9775 | -90.4877
9.9778 | -90.4886
9.9781 | -90.4891
9.9785 | -90.489
9.9788 | -90.49
9.9791 | -90.4906
9.9795 | -90.4908
9.9798 | -90.4912
9.9801 | -90.4908
9.9804 | -90.4902
9.9808 | -90.4903
9.9811 | -90.4906
9.9814 | -90.4907
9.9818 | -90.4913
9.9821 | -90.491
9.9824 | -90.4907
9.9828 | -90.491
9.9831 | -90.491
9.9834 | -90.4912
9.9837 | -90.4915
9.9841 | -90.4917
9.9844 | -90.4917
9.9847 | -90.4916
9.9851 | -90.492
9.9854 | -90.4923
9.9857 | -90.4923
9.986 | -90.4909
9.9864 | -90.4903
9.9867 | -90.4905
9.987 | -90.4909
9.9874 | -90.4912
9.9877 | -90.4916
9.988 | -90.4921
9.9883 | -90.4923
9.9887 | -90.4926
9.989 | -90.4934
9.9893 | -90.4945
9.9897 | -90.4952
9.99 | -90.4959
9.9903 | -90.4965
9.9907 | -90.4969
9.991 | -90.4971
9.9913 | -90.4973
9.9916 | -90.4974
9.992 | -90.4976
9.9923 | -90.4978
9.9926 | -90.498
9.993 | -90.4982
9.9933 | -90.4983
9.9936 | -90.4983
9.9939 | -90.499
9.9943 | -90.4997
9.9946 | -90.5004
9.9949 | -90.501
9.9953 | -90.5011
9.9956 | -90.5015
9.9959 | -90.502
9.9962 | -90.5025
9.9966 | -90.5031
9.9969 | -90.5037
9.9972 | -90.5041
9.9976 | -90.5039
9.9979 | -90.504
9.9982 | -90.5042
9.9986 | -90.5042
9.9989 | -90.5048
9.9992 | -90.5052
9.9995 | -90.5056
9.9999 | -90.5053
10.0002 | -90.5062
10.0005 | -90.5076
10.0009 | -90.5084
10.0012 | -90.509
10.0015 | -90.5096
10.0018 | -90.5099
10.0022 | -90.5102
10.0025 | -90.5105
10.0028 | -90.5108
10.0032 | -90.5115
10.0035 | -90.512
10.0038 | -90.5121
10.0041 | -90.5119
10.0045 | -90.5117
10.0048 | -90.5114
10.0051 | -90.5115
10.0055 | -90.5117
10.0058 | -90.5118
10.0061 | -90.5122
10.0065 | -90.5122
10.0068 | -90.5121
10.0071 | -90.512
10.0074 | -90.5127
10.0078 | -90.5135
10.0081 | -90.5141
10.0084 | -90.5149
10.0088 | -90.515
10.0091 | -90.5157
10.0094 | -90.5167
10.0097 | -90.5173
10.0101 | -90.5177
10.0104 | -90.5179
10.0107 | -90.5181
10.0111 | -90.5185
10.0114 | -90.5189
10.0117 | -90.5192
10.012 | -90.5197
10.0124 | -90.5203
10.0127 | -90.5209
10.013 | -90.5214
10.0134 | -90.5218
10.0137 | -90.5221
10.014 | -90.5225
10.0144 | -90.523
10.0147 | -90.5234
10.015 | -90.5234
10.0153 | -90.5236
10.0157 | -90.5239
10.016 | -90.524
10.0163 | -90.5242
10.0167 | -90.5243
10.017 | -90.5243
10.0173 | -90.5248
10.0176 | -90.5255
10.018 | -90.5258
10.0183 | -90.526
10.0186 | -90.5266
10.019 | -90.5275
10.0193 | -90.5286
10.0196 | -90.5291
10.0199 | -90.5294
10.0203 | -90.5298
10.0206 | -90.5301
10.0209 | -90.5304
10.0213 | -90.5308
10.0216 | -90.5309
10.0219 | -90.5311
10.0222 | -90.5314
10.0226 | -90.5318
10.0229 | -90.5322
10.0232 | -90.5328
10.0236 | -90.5332
10.0239 | -90.5334
10.0242 | -90.5335
10.0246 | -90.5337
10.0249 | -90.534
10.0252 | -90.5344
10.0255 | -90.5346
10.0259 | -90.5349
10.0262 | -90.5354
10.0265 | -90.5358
10.0269 | -90.5366
10.0272 | -90.5371
10.0275 | -90.5376
10.0278 | -90.5381
10.0282 | -90.5382
10.0285 | -90.5389
10.0288 | -90.5398
10.0292 | -90.5409
10.0295 | -90.5415
10.0298 | -90.5418
10.0301 | -90.5424
10.0305 | -90.5429
10.0308 | -90.5433
10.0311 | -90.5434
10.0315 | -90.5436
10.0318 | -90.544
10.0321 | -90.5446
10.0325 | -90.545
10.0328 | -90.5451
10.0331 | -90.5453
10.0334 | -90.5458
10.0338 | -90.5456
10.0341 | -90.5452
10.0344 | -90.5458
10.0348 | -90.5461
10.0351 | -90.5465
10.0354 | -90.5469
10.0357 | -90.5474
10.0361 | -90.5475
10.0364 | -90.5479
10.0367 | -90.5484
10.0371 | -90.5489
10.0374 | -90.5488
10.0377 | -90.5475
10.038 | -90.5475
10.0384 | -90.5478
10.0387 | -90.5487
10.039 | -90.5493
10.0394 | -90.5496
10.0397 | -90.5501
10.04 | -90.5504
10.0404 | -90.5509
10.0407 | -90.5514
10.041 | -90.5516
10.0413 | -90.5521
10.0417 | -90.5521
10.042 | -90.5514
10.0423 | -90.5509
10.0427 | -90.5517
10.043 | -90.5523
10.0433 | -90.5527
10.0436 | -90.553
10.044 | -90.5532
10.0443 | -90.5534
10.0446 | -90.5535
10.045 | -90.553
10.0453 | -90.5537
10.0456 | -90.5542
10.0459 | -90.5546
10.0463 | -90.5551
10.0466 | -90.5554
10.0469 | -90.5553
10.0473 | -90.5555
10.0476 | -90.556
10.0479 | -90.5564
10.0483 | -90.5567
10.0486 | -90.5568
10.0489 | -90.5569
10.0492 | -90.5571
10.0496 | -90.5576
10.0499 | -90.558
10.0502 | -90.5583
10.0506 | -90.5588
10.0509 | -90.5598
10.0512 | -90.5604
10.0515 | -90.5605
10.0519 | -90.5608
10.0522 | -90.5612
10.0525 | -90.5615
10.0529 | -90.5619
10.0532 | -90.5625
10.0535 | -90.5631
10.0538 | -90.5637
10.0542 | -90.564
10.0545 | -90.5644
10.0548 | -90.5657
10.0552 | -90.5661
10.0555 | -90.5665
10.0558 | -90.5666
10.0562 | -90.5666
10.0565 | -90.5673
10.0568 | -90.5678
10.0571 | -90.5681
10.0575 | -90.5684
10.0578 | -90.5696
10.0581 | -90.5701
10.0585 | -90.5705
10.0588 | -90.5708
10.0591 | -90.5711
10.0594 | -90.5715
10.0598 | -90.5718
10.0601 | -90.5721
10.0604 | -90.5727
10.0608 | -90.573
10.0611 | -90.5737
10.0614 | -90.5746
10.0617 | -90.5752
10.0621 | -90.5756
10.0624 | -90.5758
10.0627 | -90.576
10.0631 | -90.5764
10.0634 | -90.5767
10.0637 | -90.5771
10.064 | -90.5774
10.0644 | -90.5781
10.0647 | -90.5784
10.065 | -90.5789
10.0654 | -90.5796
10.0657 | -90.58
10.066 | -90.5807
10.0664 | -90.5812
10.0667 | -90.5818
10.067 | -90.583
10.0673 | -90.5841
10.0677 | -90.5846
10.068 | -90.5851
10.0683 | -90.5859
10.0687 | -90.5863
10.069 | -90.5868
10.0693 | -90.5874
10.0696 | -90.5879
10.07 | -90.5885
10.0703 | -90.5898
10.0706 | -90.5906
10.071 | -90.591
10.0713 | -90.5919
10.0716 | -90.5921
10.0719 | -90.5926
10.0723 | -90.593
10.0726 | -90.5938
10.0729 | -90.5951
10.0733 | -90.596
10.0736 | -90.5966
10.0739 | -90.5973
10.0743 | -90.5978
10.0746 | -90.5976
10.0749 | -90.5973
10.0752 | -90.5975
10.0756 | -90.5978
10.0759 | -90.5983
10.0762 | -90.5989
10.0766 | -90.5992
10.0769 | -90.5997
10.0772 | -90.6001
10.0775 | -90.6003
10.0779 | -90.6008
10.0782 | -90.6016
10.0785 | -90.6021
10.0789 | -90.6028
10.0792 | -90.6036
10.0795 | -90.6044
10.0798 | -90.6051
10.0802 | -90.6055
10.0805 | -90.6059
10.0808 | -90.6068
10.0812 | -90.6079
10.0815 | -90.6082
10.0818 | -90.6084
10.0822 | -90.6089
10.0825 | -90.6097
10.0828 | -90.6102
10.0831 | -90.6109
10.0835 | -90.6114
10.0838 | -90.612
10.0841 | -90.613
10.0845 | -90.6132
10.0848 | -90.6135
10.0851 | -90.6138
10.0854 | -90.6143
10.0858 | -90.6149
10.0861 | -90.6162
10.0864 | -90.6169
10.0868 | -90.6173
10.0871 | -90.6176
10.0874 | -90.6179
10.0877 | -90.6185
10.0881 | -90.6197
10.0884 | -90.62
10.0887 | -90.6205
10.0891 | -90.6206
10.0894 | -90.6208
10.0897 | -90.6213
10.0901 | -90.6217
10.0904 | -90.622
10.0907 | -90.6221
10.091 | -90.6225
10.0914 | -90.6231
10.0917 | -90.6235
10.092 | -90.6238
10.0924 | -90.624
10.0927 | -90.6244
10.093 | -90.6245
10.0933 | -90.6249
10.0937 | -90.6254
10.094 | -90.6259
10.0943 | -90.6263
10.0947 | -90.6262
10.095 | -90.6265
10.0953 | -90.6269
10.0956 | -90.6277
10.096 | -90.6286
10.0963 | -90.6291
10.0966 | -90.6308
10.097 | -90.6319
10.0973 | -90.6324
10.0976 | -90.6325
10.098 | -90.6325
10.0983 | -90.6328
10.0986 | -90.633
10.0989 | -90.6329
10.0993 | -90.6337
10.0996 | -90.6349
10.0999 | -90.6354
10.1003 | -90.6359
10.1006 | -90.6363
10.1009 | -90.6367
10.1012 | -90.6371
10.1016 | -90.6373
10.1019 | -90.6374
10.1022 | -90.6379
10.1026 | -90.6379
10.1029 | -90.6381
10.1032 | -90.6386
10.1035 | -90.639
10.1039 | -90.6394
10.1042 | -90.6401
10.1045 | -90.6404
10.1049 | -90.6409
10.1052 | -90.6411
10.1055 | -90.6415
10.1059 | -90.6417
10.1062 | -90.6422
10.1065 | -90.6427
10.1068 | -90.643
10.1072 | -90.6435
10.1075 | -90.6443
10.1078 | -90.6447
10.1082 | -90.6452
10.1085 | -90.6458
10.1088 | -90.6465
10.1091 | -90.6471
10.1095 | -90.6479
10.1098 | -90.6486
10.1101 | -90.6489
10.1105 | -90.6493
10.1108 | -90.6497
10.1111 | -90.65
10.1114 | -90.6503
10.1118 | -90.6506
10.1121 | -90.6512
10.1124 | -90.6515
10.1128 | -90.6518
10.1131 | -90.6523
10.1134 | -90.6532
10.1137 | -90.6535
10.1141 | -90.6537
10.1144 | -90.6541
10.1147 | -90.6547
10.1151 | -90.6552
10.1154 | -90.6555
10.1157 | -90.656
10.1161 | -90.6568
10.1164 | -90.657
10.1167 | -90.6569
10.117 | -90.6576
10.1174 | -90.6583
10.1177 | -90.6587
10.118 | -90.6592
10.1184 | -90.6599
10.1187 | -90.6605
10.119 | -90.6608
10.1193 | -90.6612
10.1197 | -90.6613
10.12 | -90.662
10.1203 | -90.6633
10.1207 | -90.6648
10.121 | -90.6658
10.1213 | -90.6663
10.1216 | -90.6665
10.122 | -90.6667
10.1223 | -90.6671
10.1226 | -90.6674
10.123 | -90.6681
10.1233 | -90.6685
10.1236 | -90.6689
10.124 | -90.6696
10.1243 | -90.6701
10.1246 | -90.6708
10.1249 | -90.6715
10.1253 | -90.6719
10.1256 | -90.6719
10.1259 | -90.672
10.1263 | -90.6724
10.1266 | -90.6729
10.1269 | -90.6735
10.1272 | -90.6738
10.1276 | -90.6743
10.1279 | -90.6749
10.1282 | -90.6758
10.1286 | -90.6762
10.1289 | -90.6775
10.1292 | -90.6788
10.1295 | -90.6798
10.1299 | -90.6807
10.1302 | -90.6815
10.1305 | -90.682
10.1309 | -90.6823
10.1312 | -90.6827
10.1315 | -90.6831
10.1319 | -90.6841
10.1322 | -90.6852
10.1325 | -90.6855
10.1328 | -90.6858
10.1332 | -90.6858
10.1335 | -90.6861
10.1338 | -90.6862
10.1342 | -90.6866
10.1345 | -90.687
10.1348 | -90.6876
10.1351 | -90.6886
10.1355 | -90.6892
10.1358 | -90.6897
10.1361 | -90.6903
10.1365 | -90.691
10.1368 | -90.6914
10.1371 | -90.6922
10.1374 | -90.6933
10.1378 | -90.694
10.1381 | -90.6946
10.1384 | -90.6953
10.1388 | -90.6959
10.1391 | -90.6964
10.1394 | -90.6967
10.1398 | -90.6971
10.1401 | -90.6976
10.1404 | -90.6987
10.1407 | -90.6994
10.1411 | -90.6995
10.1414 | -90.7001
10.1417 | -90.7008
10.1421 | -90.7015
10.1424 | -90.7015
10.1427 | -90.702
10.143 | -90.7031
10.1434 | -90.7032
10.1437 | -90.7036
10.144 | -90.7043
10.1444 | -90.705
10.1447 | -90.7054
10.145 | -90.7057
10.1453 | -90.7062
10.1457 | -90.7069
10.146 | -90.7075
10.1463 | -90.7077
10.1467 | -90.7079
10.147 | -90.7085
10.1473 | -90.7091
10.1477 | -90.7096
10.148 | -90.7098
10.1483 | -90.7105
10.1486 | -90.711
10.149 | -90.7115
10.1493 | -90.7125
10.1496 | -90.7135
10.15 | -90.7138
10.1503 | -90.7141
10.1506 | -90.7149
10.1509 | -90.7157
10.1513 | -90.7166
10.1516 | -90.7171
10.1519 | -90.7179
10.1523 | -90.7187
10.1526 | -90.7186
10.1529 | -90.7187
10.1532 | -90.7193
10.1536 | -90.72
10.1539 | -90.7202
10.1542 | -90.7207
10.1546 | -90.7216
10.1549 | -90.7218
10.1552 | -90.7222
10.1556 | -90.7228
10.1559 | -90.7238
10.1562 | -90.725
10.1565 | -90.7255
10.1569 | -90.7259
10.1572 | -90.7266
10.1575 | -90.7273
10.1579 | -90.7283
10.1582 | -90.7289
10.1585 | -90.7294
10.1588 | -90.73
10.1592 | -90.7305
10.1595 | -90.7311
10.1598 | -90.7313
10.1602 | -90.7307
10.1605 | -90.731
10.1608 | -90.7313
10.1611 | -90.7316
10.1615 | -90.7325
10.1618 | -90.7332
10.1621 | -90.7345
10.1625 | -90.7348
10.1628 | -90.7351
10.1631 | -90.7356
10.1634 | -90.736
10.1638 | -90.7365
10.1641 | -90.7369
10.1644 | -90.7375
10.1648 | -90.7379
10.1651 | -90.7387
10.1654 | -90.7405
10.1658 | -90.7415
10.1661 | -90.7422
10.1664 | -90.7426
10.1667 | -90.7428
10.1671 | -90.7433
10.1674 | -90.7438
10.1677 | -90.7438
10.1681 | -90.744
10.1684 | -90.7446
10.1687 | -90.7449
10.169 | -90.7448
10.1694 | -90.7443
10.1697 | -90.7443
10.17 | -90.7445
10.1704 | -90.7446
10.1707 | -90.7443
10.171 | -90.7443
10.1713 | -90.745
10.1717 | -90.7455
10.172 | -90.746
10.1723 | -90.7468
10.1727 | -90.7474
10.173 | -90.7487
10.1733 | -90.7496
10.1737 | -90.75
10.174 | -90.7501
10.1743 | -90.7502
10.1746 | -90.7504
10.175 | -90.7509
10.1753 | -90.7515
10.1756 | -90.7521
10.176 | -90.7527
10.1763 | -90.7531
10.1766 | -90.7538
10.1769 | -90.7547
10.1773 | -90.7552
10.1776 | -90.7553
10.1779 | -90.7554
10.1783 | -90.7557
10.1786 | -90.7567
10.1789 | -90.758
10.1792 | -90.7585
10.1796 | -90.7584
10.1799 | -90.7586
10.1802 | -90.7591
10.1806 | -90.7597
10.1809 | -90.7606
10.1812 | -90.7616
10.1816 | -90.7625
10.1819 | -90.7631
10.1822 | -90.7639
10.1825 | -90.7648
10.1829 | -90.7653
10.1832 | -90.7658
10.1835 | -90.7667
10.1839 | -90.7671
10.1842 | -90.7673
10.1845 | -90.7678
10.1848 | -90.7685
10.1852 | -90.769
10.1855 | -90.7695
10.1858 | -90.7699
10.1862 | -90.7707
10.1865 | -90.7717
10.1868 | -90.7725
10.1871 | -90.7729
10.1875 | -90.7727
10.1878 | -90.7726
10.1881 | -90.7725
10.1885 | -90.7723
10.1888 | -90.7723
10.1891 | -90.7728
10.1895 | -90.7737
10.1898 | -90.7748
10.1901 | -90.7757
10.1904 | -90.7762
10.1908 | -90.777
10.1911 | -90.7776
10.1914 | -90.7781
10.1918 | -90.7784
10.1921 | -90.7786
10.1924 | -90.779
10.1927 | -90.7794
10.1931 | -90.7803
10.1934 | -90.7808
10.1937 | -90.7812
10.1941 | -90.7816
10.1944 | -90.7824
10.1947 | -90.7828
10.195 | -90.7835
10.1954 | -90.7838
10.1957 | -90.7844
10.196 | -90.7848
10.1964 | -90.7851
10.1967 | -90.7858
10.197 | -90.7865
10.1974 | -90.7869
10.1977 | -90.7878
10.198 | -90.7884
10.1983 | -90.7894
10.1987 | -90.7903
10.199 | -90.7913
10.1993 | -90.7921
10.1997 | -90.7926
10.2 | -90.7934
10.2003 | -90.7944
10.2006 | -90.7947
10.201 | -90.7949
10.2013 | -90.7957
10.2016 | -90.7965
10.202 | -90.797
10.2023 | -90.7975
10.2026 | -90.7982
10.2029 | -90.799
10.2033 | -90.7999
10.2036 | -90.8004
10.2039 | -90.8008
10.2043 | -90.8015
10.2046 | -90.8017
10.2049 | -90.8021
10.2052 | -90.8027
10.2056 | -90.8035
10.2059 | -90.804
10.2062 | -90.8046
10.2066 | -90.8051
10.2069 | -90.8053
10.2072 | -90.8055
10.2076 | -90.8062
10.2079 | -90.8069
10.2082 | -90.8073
10.2085 | -90.808
10.2089 | -90.8087
10.2092 | -90.8091
10.2095 | -90.8096
10.2099 | -90.8101
10.2102 | -90.8108
10.2105 | -90.8112
10.2108 | -90.8116
10.2112 | -90.812
10.2115 | -90.8124
10.2118 | -90.813
10.2122 | -90.8139
10.2125 | -90.8148
10.2128 | -90.8157
10.2132 | -90.8164
10.2135 | -90.8169
10.2138 | -90.8174
10.2141 | -90.8178
10.2145 | -90.8182
10.2148 | -90.8188
10.2151 | -90.8194
10.2155 | -90.8199
10.2158 | -90.8203
10.2161 | -90.821
10.2164 | -90.8217
10.2168 | -90.822
10.2171 | -90.8222
10.2174 | -90.8226
10.2178 | -90.8233
10.2181 | -90.8243
10.2184 | -90.825
10.2187 | -90.8256
10.2191 | -90.8262
10.2194 | -90.8268
10.2197 | -90.8273
10.2201 | -90.8274
10.2204 | -90.8276
10.2207 | -90.8282
10.221 | -90.829
10.2214 | -90.8298
10.2217 | -90.8305
10.222 | -90.8307
10.2224 | -90.831
10.2227 | -90.8316
10.223 | -90.8321
10.2234 | -90.8327
10.2237 | -90.8332
10.224 | -90.8341
10.2243 | -90.8347
10.2247 | -90.8353
10.225 | -90.8361
10.2253 | -90.8366
10.2257 | -90.8372
10.226 | -90.8378
10.2263 | -90.8379
10.2266 | -90.8373
10.227 | -90.8366
10.2273 | -90.837
10.2276 | -90.8378
10.228 | -90.8383
10.2283 | -90.8385
10.2286 | -90.8389
10.2289 | -90.8392
10.2293 | -90.8396
10.2296 | -90.8407
10.2299 | -90.8421
10.2303 | -90.8431
10.2306 | -90.8437
10.2309 | -90.8442
10.2313 | -90.8445
10.2316 | -90.8448
10.2319 | -90.8452
10.2322 | -90.8455
10.2326 | -90.8467
10.2329 | -90.8483
10.2332 | -90.8492
10.2336 | -90.8497
10.2339 | -90.8506
10.2342 | -90.8512
10.2345 | -90.8514
10.2349 | -90.8521
10.2352 | -90.8527
10.2355 | -90.8535
10.2359 | -90.854
10.2362 | -90.8539
10.2365 | -90.8545
10.2368 | -90.8553
10.2372 | -90.8563
10.2375 | -90.8574
10.2378 | -90.8578
10.2382 | -90.8579
10.2385 | -90.8584
10.2388 | -90.859
10.2392 | -90.8593
10.2395 | -90.8597
10.2398 | -90.8602
10.2401 | -90.8608
10.2405 | -90.8608
10.2408 | -90.8607
10.2411 | -90.8613
10.2415 | -90.862
10.2418 | -90.8625
10.2421 | -90.863
10.2424 | -90.8636
10.2428 | -90.8642
10.2431 | -90.8646
10.2434 | -90.8648
10.2438 | -90.8654
10.2441 | -90.8657
10.2444 | -90.8661
10.2447 | -90.8667
10.2451 | -90.8672
10.2454 | -90.8678
10.2457 | -90.8684
10.2461 | -90.8693
10.2464 | -90.8699
10.2467 | -90.8704
10.2471 | -90.8714
10.2474 | -90.8723
10.2477 | -90.8732
10.248 | -90.874
10.2484 | -90.8742
10.2487 | -90.8748
10.249 | -90.8754
10.2494 | -90.8759
10.2497 | -90.8763
10.25 | -90.8765
10.2503 | -90.8772
10.2507 | -90.8777
10.251 | -90.8783
10.2513 | -90.8784
10.2517 | -90.8788
10.252 | -90.8799
10.2523 | -90.881
10.2526 | -90.8816
10.253 | -90.8821
10.2533 | -90.8827
10.2536 | -90.8837
10.254 | -90.8857
10.2543 | -90.8864
10.2546 | -90.8869
10.2549 | -90.8875
10.2553 | -90.888
10.2556 | -90.8884
10.2559 | -90.8892
10.2563 | -90.8896
10.2566 | -90.8899
10.2569 | -90.8906
10.2573 | -90.8911
10.2576 | -90.8914
10.2579 | -90.8916
10.2582 | -90.8925
10.2586 | -90.8929
10.2589 | -90.8941
10.2592 | -90.8954
10.2596 | -90.8959
10.2599 | -90.8966
10.2602 | -90.8974
10.2605 | -90.8981
10.2609 | -90.8986
10.2612 | -90.8993
10.2615 | -90.8999
10.2619 | -90.9005
10.2622 | -90.9013
10.2625 | -90.9011
10.2629 | -90.9009
10.2632 | -90.9013
10.2635 | -90.9014
10.2638 | -90.9008
10.2642 | -90.9009
10.2645 | -90.9017
10.2648 | -90.903
10.2652 | -90.9035
10.2655 | -90.9039
10.2658 | -90.9039
10.2661 | -90.9048
10.2665 | -90.9056
10.2668 | -90.9057
10.2671 | -90.9049
10.2675 | -90.9055
10.2678 | -90.9061
10.2681 | -90.9066
10.2684 | -90.9071
10.2688 | -90.908
10.2691 | -90.9088
10.2694 | -90.9094
10.2698 | -90.91
10.2701 | -90.9109
10.2704 | -90.9121
10.2707 | -90.9122
10.2711 | -90.9118
10.2714 | -90.9122
10.2717 | -90.9127
10.2721 | -90.9134
10.2724 | -90.9144
10.2727 | -90.9155
10.2731 | -90.916
10.2734 | -90.9166
10.2737 | -90.9173
10.274 | -90.9182
10.2744 | -90.9195
10.2747 | -90.9202
10.275 | -90.9209
10.2754 | -90.9218
10.2757 | -90.9225
10.276 | -90.9232
10.2763 | -90.9238
10.2767 | -90.9243
10.277 | -90.9249
10.2773 | -90.9253
10.2777 | -90.9256
10.278 | -90.9264
10.2783 | -90.9271
10.2786 | -90.9277
10.279 | -90.9282
10.2793 | -90.9287
10.2796 | -90.9293
10.28 | -90.9298
10.2803 | -90.9302
10.2806 | -90.9312
10.281 | -90.932
10.2813 | -90.9325
10.2816 | -90.9335
10.2819 | -90.9341
10.2823 | -90.9349
10.2826 | -90.9352
10.2829 | -90.9356
10.2833 | -90.9362
10.2836 | -90.9369
10.2839 | -90.9376
10.2842 | -90.9382
10.2846 | -90.939
10.2849 | -90.9395
10.2852 | -90.9402
10.2856 | -90.9407
10.2859 | -90.9412
10.2862 | -90.9417
10.2865 | -90.9422
10.2869 | -90.9431
10.2872 | -90.945
10.2875 | -90.9458
10.2879 | -90.9465
10.2882 | -90.9474
10.2885 | -90.9484
10.2889 | -90.9493
10.2892 | -90.9498
10.2895 | -90.9503
10.2898 | -90.9509
10.2902 | -90.9516
10.2905 | -90.9513
10.2908 | -90.9518
10.2912 | -90.9535
10.2915 | -90.9541
10.2918 | -90.9544
10.2921 | -90.955
10.2925 | -90.9555
10.2928 | -90.9559
10.2931 | -90.9565
10.2935 | -90.9578
10.2938 | -90.9592
10.2941 | -90.9597
10.2944 | -90.9602
10.2948 | -90.9612
10.2951 | -90.9618
10.2954 | -90.962
10.2958 | -90.9625
10.2961 | -90.9633
10.2964 | -90.9634
10.2968 | -90.9636
10.2971 | -90.9644
10.2974 | -90.9651
10.2977 | -90.9657
10.2981 | -90.9663
10.2984 | -90.9671
10.2987 | -90.9678
10.2991 | -90.9683
10.2994 | -90.9689
10.2997 | -90.9691
10.3 | -90.9692
10.3004 | -90.9704
10.3007 | -90.9716
10.301 | -90.9719
10.3014 | -90.972
10.3017 | -90.9725
10.302 | -90.9731
10.3023 | -90.9741
10.3027 | -90.975
10.303 | -90.9757
10.3033 | -90.9748
10.3037 | -90.975
10.304 | -90.9758
10.3043 | -90.9766
10.3047 | -90.9772
10.305 | -90.9777
10.3053 | -90.9784
10.3056 | -90.9794
10.306 | -90.9799
10.3063 | -90.981
10.3066 | -90.9822
10.307 | -90.983
10.3073 | -90.9835
10.3076 | -90.9841
10.3079 | -90.9849
10.3083 | -90.9852
10.3086 | -90.9854
10.3089 | -90.9864
10.3093 | -90.9872
10.3096 | -90.9877
10.3099 | -90.9885
10.3102 | -90.9908
10.3106 | -90.9931
10.3109 | -90.9941
10.3112 | -90.9947
10.3116 | -90.9953
10.3119 | -90.9963
10.3122 | -90.997
10.3125 | -90.9973
10.3129 | -90.9978
10.3132 | -90.9984
10.3135 | -90.999
10.3139 | -90.9998
10.3142 | -91.0014
10.3145 | -91.003
10.3149 | -91.0035
10.3152 | -91.0043
10.3155 | -91.0048
10.3158 | -91.0056
10.3162 | -91.0057
10.3165 | -91.0058
10.3168 | -91.0063
10.3172 | -91.0068
10.3175 | -91.0071
10.3178 | -91.0075
10.3181 | -91.008
10.3185 | -91.0087
10.3188 | -91.0094
10.3191 | -91.01
10.3195 | -91.0109
10.3198 | -91.0118
10.3201 | -91.0126
10.3204 | -91.0135
10.3208 | -91.0144
10.3211 | -91.0151
10.3214 | -91.0156
10.3218 | -91.016
10.3221 | -91.0171
10.3224 | -91.0179
10.3228 | -91.0189
10.3231 | -91.0193
10.3234 | -91.0197
10.3237 | -91.0201
10.3241 | -91.0202
10.3244 | -91.0202
10.3247 | -91.0206
10.3251 | -91.0212
10.3254 | -91.0217
10.3257 | -91.0226
10.326 | -91.0234
10.3264 | -91.0241
10.3267 | -91.0245
10.327 | -91.0247
10.3274 | -91.0256
10.3277 | -91.0257
10.328 | -91.0259
10.3283 | -91.027
10.3287 | -91.0274
10.329 | -91.0276
10.3293 | -91.0282
10.3297 | -91.0294
10.33 | -91.0301
10.3303 | -91.0306
10.3307 | -91.0311
10.331 | -91.0318
10.3313 | -91.0323
10.3316 | -91.0324
10.332 | -91.0322
10.3323 | -91.0327
10.3326 | -91.0334
10.333 | -91.0337
10.3333 | -91.0341
10.3336 | -91.0349
10.3339 | -91.0358
10.3343 | -91.0365
10.3346 | -91.0375
10.3349 | -91.0379
10.3353 | -91.0384
10.3356 | -91.039
10.3359 | -91.0389
10.3362 | -91.0394
10.3366 | -91.0401
10.3369 | -91.0407
10.3372 | -91.0405
10.3376 | -91.0408
10.3379 | -91.041
10.3382 | -91.0418
10.3386 | -91.0424
10.3389 | -91.0431
10.3392 | -91.0439
10.3395 | -91.0445
10.3399 | -91.0452
10.3402 | -91.0459
10.3405 | -91.0468
10.3409 | -91.0476
10.3412 | -91.0482
10.3415 | -91.0488
10.3418 | -91.0499
10.3422 | -91.0504
10.3425 | -91.0514
10.3428 | -91.0529
10.3432 | -91.0535
10.3435 | -91.0537
10.3438 | -91.0542
10.3441 | -91.0548
10.3445 | -91.0552
10.3448 | -91.0554
10.3451 | -91.056
10.3455 | -91.0569
10.3458 | -91.0577
10.3461 | -91.0586
10.3465 | -91.0597
10.3468 | -91.0606
10.3471 | -91.0608
10.3474 | -91.0607
10.3478 | -91.061
10.3481 | -91.0614
10.3484 | -91.0618
10.3488 | -91.0624
10.3491 | -91.0635
10.3494 | -91.0648
10.3497 | -91.0655
10.3501 | -91.0653
10.3504 | -91.065
10.3507 | -91.0654
10.3511 | -91.066
10.3514 | -91.0664
10.3517 | -91.0668
10.352 | -91.0676
10.3524 | -91.0683
10.3527 | -91.0689
10.353 | -91.0695
10.3534 | -91.0699
10.3537 | -91.0701
10.354 | -91.0706
10.3544 | -91.0709
10.3547 | -91.0708
10.355 | -91.0712
10.3553 | -91.0719
10.3557 | -91.0726
10.356 | -91.0739
10.3563 | -91.0749
10.3567 | -91.0756
10.357 | -91.0759
10.3573 | -91.0759
10.3576 | -91.0762
10.358 | -91.0766
10.3583 | -91.0771
10.3586 | -91.0776
10.359 | -91.0781
10.3593 | -91.0788
10.3596 | -91.0783
10.3599 | -91.0781
10.3603 | -91.0782
10.3606 | -91.0789
10.3609 | -91.0791
10.3613 | -91.0796
10.3616 | -91.08
10.3619 | -91.0803
10.3622 | -91.0803
10.3626 | -91.0803
10.3629 | -91.0812
10.3632 | -91.0818
10.3636 | -91.0825
10.3639 | -91.0833
10.3642 | -91.0839
10.3646 | -91.0844
10.3649 | -91.0847
10.3652 | -91.0852
10.3655 | -91.0858
10.3659 | -91.0856
10.3662 | -91.0857
10.3665 | -91.0863
10.3669 | -91.0874
10.3672 | -91.0884
10.3675 | -91.0889
10.3678 | -91.0896
10.3682 | -91.0901
10.3685 | -91.0901
10.3688 | -91.0905
10.3692 | -91.091
10.3695 | -91.0919
10.3698 | -91.0925
10.3701 | -91.0929
10.3705 | -91.0935
10.3708 | -91.0941
10.3711 | -91.0944
10.3715 | -91.0948
10.3718 | -91.0953
10.3721 | -91.0958
10.3725 | -91.0962
10.3728 | -91.0966
10.3731 | -91.0971
10.3734 | -91.0972
10.3738 | -91.0974
10.3741 | -91.0979
10.3744 | -91.0985
10.3748 | -91.099
10.3751 | -91.0994
10.3754 | -91.0992
10.3757 | -91.0993
10.3761 | -91.0997
10.3764 | -91.1001
10.3767 | -91.1007
10.3771 | -91.1014
10.3774 | -91.1019
10.3777 | -91.1023
10.378 | -91.1028
10.3784 | -91.1027
10.3787 | -91.1026
10.379 | -91.1027
10.3794 | -91.1025
10.3797 | -91.1025
10.38 | -91.103
10.3804 | -91.1049
10.3807 | -91.1067
10.381 | -91.1073
10.3813 | -91.108
10.3817 | -91.1087
10.382 | -91.1091
10.3823 | -91.1099
10.3827 | -91.1105
10.383 | -91.1111
10.3833 | -91.1115
10.3836 | -91.112
10.384 | -91.1122
10.3843 | -91.1122
10.3846 | -91.1125
10.385 | -91.113
10.3853 | -91.1133
10.3856 | -91.1136
10.3859 | -91.1136
10.3863 | -91.1133
10.3866 | -91.1138
10.3869 | -91.114
10.3873 | -91.1142
10.3876 | -91.115
10.3879 | -91.1158
10.3883 | -91.116
10.3886 | -91.1165
10.3889 | -91.1172
10.3892 | -91.1175
10.3896 | -91.1179
10.3899 | -91.1179
10.3902 | -91.118
10.3906 | -91.1186
10.3909 | -91.1194
10.3912 | -91.1199
10.3915 | -91.1202
10.3919 | -91.1206
10.3922 | -91.1209
10.3925 | -91.1215
10.3929 | -91.1228
10.3932 | -91.1236
10.3935 | -91.1245
10.3938 | -91.1258
10.3942 | -91.1263
10.3945 | -91.1268
10.3948 | -91.1274
10.3952 | -91.1286
10.3955 | -91.1293
10.3958 | -91.1297
10.3962 | -91.1308
10.3965 | -91.1325
10.3968 | -91.1333
10.3971 | -91.1337
10.3975 | -91.1342
10.3978 | -91.1347
10.3981 | -91.135
10.3985 | -91.1354
10.3988 | -91.1358
10.3991 | -91.1362
10.3994 | -91.1364
10.3998 | -91.1368
10.4001 | -91.1379
10.4004 | -91.1389
10.4008 | -91.1396
10.4011 | -91.1402
10.4014 | -91.1406
10.4017 | -91.1404
10.4021 | -91.1415
10.4024 | -91.1422
10.4027 | -91.1426
10.4031 | -91.1432
10.4034 | -91.144
10.4037 | -91.1443
10.404 | -91.1447
10.4044 | -91.1448
10.4047 | -91.1454
10.405 | -91.1471
10.4054 | -91.1482
10.4057 | -91.1485
10.406 | -91.1488
10.4064 | -91.148
10.4067 | -91.1476
10.407 | -91.1477
10.4073 | -91.1478
10.4077 | -91.1488
10.408 | -91.1495
10.4083 | -91.1505
10.4087 | -91.1517
10.409 | -91.1521
10.4093 | -91.1521
10.4096 | -91.1523
10.41 | -91.1526
10.4103 | -91.1531
10.4106 | -91.1538
10.411 | -91.1539
10.4113 | -91.1545
10.4116 | -91.1554
10.4119 | -91.1557
10.4123 | -91.156
10.4126 | -91.156
10.4129 | -91.156
10.4133 | -91.1563
10.4136 | -91.1569
10.4139 | -91.1576
10.4143 | -91.1584
10.4146 | -91.1584
10.4149 | -91.1584
10.4152 | -91.1587
10.4156 | -91.159
10.4159 | -91.1593
10.4162 | -91.1598
10.4166 | -91.1601
10.4169 | -91.1604
10.4172 | -91.1611
10.4175 | -91.1615
10.4179 | -91.1625
10.4182 | -91.1628
10.4185 | -91.1633
10.4189 | -91.1641
10.4192 | -91.1647
10.4195 | -91.1651
10.4198 | -91.1655
10.4202 | -91.1657
10.4205 | -91.1662
10.4208 | -91.1667
10.4212 | -91.1669
10.4215 | -91.1671
10.4218 | -91.1667
10.4222 | -91.1665
10.4225 | -91.166
10.4228 | -91.1658
10.4231 | -91.1662
10.4235 | -91.1664
10.4238 | -91.1666
10.4241 | -91.1666
10.4245 | -91.1665
10.4248 | -91.1666
10.4251 | -91.1669
10.4254 | -91.1672
10.4258 | -91.1676
10.4261 | -91.1679
10.4264 | -91.1683
10.4268 | -91.1688
10.4271 | -91.1691
10.4274 | -91.1692
10.4277 | -91.1697
10.4281 | -91.1699
10.4284 | -91.1701
10.4287 | -91.1706
10.4291 | -91.1705
10.4294 | -91.1706
10.4297 | -91.1713
10.4301 | -91.1721
10.4304 | -91.1722
10.4307 | -91.172
10.431 | -91.172
10.4314 | -91.1724
10.4317 | -91.1725
10.432 | -91.1727
10.4324 | -91.1727
10.4327 | -91.173
10.433 | -91.1733
10.4333 | -91.1737
10.4337 | -91.1742
10.434 | -91.175
10.4343 | -91.1753
10.4347 | -91.1753
10.435 | -91.175
10.4353 | -91.1751
10.4356 | -91.1758
10.436 | -91.1761
10.4363 | -91.1766
10.4366 | -91.177
10.437 | -91.1774
10.4373 | -91.1778
10.4376 | -91.178
10.438 | -91.1786
10.4383 | -91.179
10.4386 | -91.1794
10.4389 | -91.1802
10.4393 | -91.1805
10.4396 | -91.1806
10.4399 | -91.1807
10.4403 | -91.181
10.4406 | -91.1812
10.4409 | -91.1814
10.4412 | -91.1817
10.4416 | -91.1819
10.4419 | -91.1822
10.4422 | -91.1826
10.4426 | -91.1832
10.4429 | -91.1839
10.4432 | -91.1845
10.4435 | -91.185
10.4439 | -91.1855
10.4442 | -91.1858
10.4445 | -91.1863
10.4449 | -91.1867
10.4452 | -91.187
10.4455 | -91.1874
10.4459 | -91.1876
10.4462 | -91.1886
10.4465 | -91.1888
10.4468 | -91.1895
10.4472 | -91.1899
10.4475 | -91.19
10.4478 | -91.1904
10.4482 | -91.1907
10.4485 | -91.1915
10.4488 | -91.1916
10.4491 | -91.1918
10.4495 | -91.1925
10.4498 | -91.1931
10.4501 | -91.1932
10.4505 | -91.1933
10.4508 | -91.1936
10.4511 | -91.1937
10.4514 | -91.1942
10.4518 | -91.195
10.4521 | -91.196
10.4524 | -91.1963
10.4528 | -91.1959
10.4531 | -91.1962
10.4534 | -91.1968
10.4537 | -91.1971
10.4541 | -91.1974
10.4544 | -91.1974
10.4547 | -91.1975
10.4551 | -91.1983
10.4554 | -91.1979
10.4557 | -91.1989
10.4561 | -91.1993
10.4564 | -91.1995
10.4567 | -91.2001
10.457 | -91.2006
10.4574 | -91.2014
10.4577 | -91.2016
10.458 | -91.2016
10.4584 | -91.2019
10.4587 | -91.2023
10.459 | -91.2025
10.4593 | -91.2029
10.4597 | -91.2034
10.46 | -91.2036
10.4603 | -91.2038
10.4607 | -91.204
10.461 | -91.2041
10.4613 | -91.2043
10.4616 | -91.2044
10.462 | -91.205
10.4623 | -91.2059
10.4626 | -91.2063
10.463 | -91.2063
10.4633 | -91.2068
10.4636 | -91.2072
10.464 | -91.2069
10.4643 | -91.2074
10.4646 | -91.2081
10.4649 | -91.2084
10.4653 | -91.2083
10.4656 | -91.2079
10.4659 | -91.2086
10.4663 | -91.2092
10.4666 | -91.2094
10.4669 | -91.2096
10.4672 | -91.21
10.4676 | -91.2099
10.4679 | -91.2101
10.4682 | -91.2103
10.4686 | -91.2111
10.4689 | -91.2115
10.4692 | -91.2117
10.4695 | -91.2118
10.4699 | -91.2117
10.4702 | -91.2117
10.4705 | -91.2119
10.4709 | -91.2121
10.4712 | -91.2124
10.4715 | -91.2121
10.4719 | -91.2118
10.4722 | -91.2119
10.4725 | -91.212
10.4728 | -91.2118
10.4732 | -91.2122
10.4735 | -91.2124
10.4738 | -91.2126
10.4742 | -91.2128
10.4745 | -91.2131
10.4748 | -91.2133
10.4751 | -91.2132
10.4755 | -91.2132
10.4758 | -91.213
10.4761 | -91.213
10.4765 | -91.2132
10.4768 | -91.2135
10.4771 | -91.2139
10.4774 | -91.2138
10.4778 | -91.2139
10.4781 | -91.2147
10.4784 | -91.2156
10.4788 | -91.2158
10.4791 | -91.2159
10.4794 | -91.2164
10.4798 | -91.2165
10.4801 | -91.2165
10.4804 | -91.2171
10.4807 | -91.2175
10.4811 | -91.2175
10.4814 | -91.2177
10.4817 | -91.2176
10.4821 | -91.2176
10.4824 | -91.2179
10.4827 | -91.2183
10.483 | -91.2191
10.4834 | -91.2199
10.4837 | -91.2202
10.484 | -91.2205
10.4844 | -91.2207
10.4847 | -91.2211
10.485 | -91.2215
10.4853 | -91.2218
10.4857 | -91.2219
10.486 | -91.2223
10.4863 | -91.2226
10.4867 | -91.223
10.487 | -91.2236
10.4873 | -91.2241
10.4877 | -91.2243
10.488 | -91.2245
10.4883 | -91.2249
10.4886 | -91.2248
10.489 | -91.2248
10.4893 | -91.2252
10.4896 | -91.2258
10.49 | -91.2261
10.4903 | -91.2264
10.4906 | -91.2269
10.4909 | -91.2276
10.4913 | -91.2284
10.4916 | -91.2286
10.4919 | -91.229
10.4923 | -91.2294
10.4926 | -91.2295
10.4929 | -91.2295
10.4932 | -91.2296
10.4936 | -91.2297
10.4939 | -91.23
10.4942 | -91.2302
10.4946 | -91.2306
10.4949 | -91.2308
10.4952 | -91.2311
10.4955 | -91.2313
10.4959 | -91.2315
10.4962 | -91.2318
10.4965 | -91.2325
10.4969 | -91.2333
10.4972 | -91.2339
10.4975 | -91.2343
10.4979 | -91.2346
10.4982 | -91.2347
10.4985 | -91.2349
10.4988 | -91.2349
10.4992 | -91.2353
10.4995 | -91.2356
10.4998 | -91.2358
10.5002 | -91.236
10.5005 | -91.2358
10.5008 | -91.236
10.5011 | -91.2359
10.5015 | -91.2359
10.5018 | -91.2368
10.5021 | -91.2374
10.5025 | -91.2379
10.5028 | -91.2376
10.5031 | -91.2381
10.5034 | -91.2392
10.5038 | -91.2409
10.5041 | -91.2411
10.5044 | -91.2416
10.5048 | -91.2423
10.5051 | -91.2429
10.5054 | -91.2437
10.5058 | -91.2436
10.5061 | -91.2435
10.5064 | -91.2435
10.5067 | -91.2438
10.5071 | -91.2442
10.5074 | -91.2448
10.5077 | -91.2457
10.5081 | -91.2461
10.5084 | -91.2462
10.5087 | -91.2464
10.509 | -91.2467
10.5094 | -91.2469
10.5097 | -91.247
10.51 | -91.2473
10.5104 | -91.2478
10.5107 | -91.2481
10.511 | -91.2472
10.5113 | -91.2466
10.5117 | -91.2472
10.512 | -91.2468
10.5123 | -91.2467
10.5127 | -91.2469
10.513 | -91.247
10.5133 | -91.2474
10.5137 | -91.2478
10.514 | -91.2478
10.5143 | -91.2478
10.5146 | -91.2485
10.515 | -91.2493
10.5153 | -91.2499
10.5156 | -91.2502
10.516 | -91.2503
10.5163 | -91.2503
10.5166 | -91.25
10.5169 | -91.2498
10.5173 | -91.25
10.5176 | -91.2499
10.5179 | -91.25
10.5183 | -91.2502
10.5186 | -91.2502
10.5189 | -91.2505
10.5192 | -91.2507
10.5196 | -91.2507
10.5199 | -91.2506
10.5202 | -91.2505
10.5206 | -91.2505
10.5209 | -91.2511
10.5212 | -91.251
10.5216 | -91.2514
10.5219 | -91.251
10.5222 | -91.251
10.5225 | -91.251
10.5229 | -91.251
10.5232 | -91.2507
10.5235 | -91.2513
10.5239 | -91.2519
10.5242 | -91.252
10.5245 | -91.2519
10.5248 | -91.2522
10.5252 | -91.2523
10.5255 | -91.2524
10.5258 | -91.2527
10.5262 | -91.2531
10.5265 | -91.2534
10.5268 | -91.2535
10.5271 | -91.2543
10.5275 | -91.2548
10.5278 | -91.2547
10.5281 | -91.2546
10.5285 | -91.2546
10.5288 | -91.2548
10.5291 | -91.2552
10.5295 | -91.2553
10.5298 | -91.2557
10.5301 | -91.2558
10.5304 | -91.2563
10.5308 | -91.2567
10.5311 | -91.2567
10.5314 | -91.2571
10.5318 | -91.2581
10.5321 | -91.2583
10.5324 | -91.2582
10.5327 | -91.2579
10.5331 | -91.2578
10.5334 | -91.258
10.5337 | -91.2581
10.5341 | -91.2584
10.5344 | -91.2587
10.5347 | -91.2586
10.535 | -91.2581
10.5354 | -91.2579
10.5357 | -91.258
10.536 | -91.2582
10.5364 | -91.2585
10.5367 | -91.2585
10.537 | -91.2587
10.5374 | -91.2589
10.5377 | -91.2592
10.538 | -91.2596
10.5383 | -91.2599
10.5387 | -91.26
10.539 | -91.2599
10.5393 | -91.2592
10.5397 | -91.2589
10.54 | -91.2591
10.5403 | -91.2593
10.5406 | -91.2599
10.541 | -91.2602
10.5413 | -91.2604
10.5416 | -91.2607
10.542 | -91.2607
10.5423 | -91.2606
10.5426 | -91.2602
10.5429 | -91.26
10.5433 | -91.26
10.5436 | -91.2604
10.5439 | -91.2604
10.5443 | -91.2601
10.5446 | -91.26
10.5449 | -91.26
10.5452 | -91.261
10.5456 | -91.2611
10.5459 | -91.261
10.5462 | -91.2611
10.5466 | -91.261
10.5469 | -91.261
10.5472 | -91.2612
10.5476 | -91.2614
10.5479 | -91.2616
10.5482 | -91.2616
10.5485 | -91.2617
10.5489 | -91.2618
10.5492 | -91.2621
10.5495 | -91.2623
10.5499 | -91.2623
10.5502 | -91.2624
10.5505 | -91.2628
10.5508 | -91.263
10.5512 | -91.2632
10.5515 | -91.2635
10.5518 | -91.2639
10.5522 | -91.2641
10.5525 | -91.2644
10.5528 | -91.2645
10.5531 | -91.2647
10.5535 | -91.2649
10.5538 | -91.2652
10.5541 | -91.2653
10.5545 | -91.2653
10.5548 | -91.2656
10.5551 | -91.2665
10.5555 | -91.2678
10.5558 | -91.2683
10.5561 | -91.2682
10.5564 | -91.2682
10.5568 | -91.2682
10.5571 | -91.2682
10.5574 | -91.2681
10.5578 | -91.268
10.5581 | -91.2681
10.5584 | -91.2682
10.5587 | -91.2685
10.5591 | -91.269
10.5594 | -91.2692
10.5597 | -91.2694
10.5601 | -91.2694
10.5604 | -91.2693
10.5607 | -91.2694
10.561 | -91.2695
10.5614 | -91.2698
10.5617 | -91.2703
10.562 | -91.2708
10.5624 | -91.2709
10.5627 | -91.2709
10.563 | -91.2703
10.5634 | -91.2699
10.5637 | -91.2699
10.564 | -91.27
10.5643 | -91.2704
10.5647 | -91.2711
10.565 | -91.2719
10.5653 | -91.2721
10.5657 | -91.2721
10.566 | -91.2726
10.5663 | -91.2736
10.5666 | -91.2745
10.567 | -91.2751
10.5673 | -91.2754
10.5676 | -91.2755
10.568 | -91.2757
10.5683 | -91.2758
10.5686 | -91.276
10.5689 | -91.2758
10.5693 | -91.2758
10.5696 | -91.276
10.5699 | -91.2762
10.5703 | -91.2764
10.5706 | -91.2766
10.5709 | -91.2771
10.5713 | -91.2774
10.5716 | -91.2771
10.5719 | -91.2773
10.5722 | -91.2773
10.5726 | -91.2774
10.5729 | -91.2775
10.5732 | -91.2775
10.5736 | -91.2775
10.5739 | -91.2777
10.5742 | -91.278
10.5745 | -91.2778
10.5749 | -91.278
10.5752 | -91.278
10.5755 | -91.2779
10.5759 | -91.2782
10.5762 | -91.2786
10.5765 | -91.2788
10.5768 | -91.279
10.5772 | -91.2791
10.5775 | -91.2794
10.5778 | -91.2797
10.5782 | -91.2802
10.5785 | -91.2805
10.5788 | -91.2806
10.5792 | -91.2807
10.5795 | -91.2809
10.5798 | -91.281
10.5801 | -91.2811
10.5805 | -91.2817
10.5808 | -91.2819
10.5811 | -91.2818
10.5815 | -91.2824
10.5818 | -91.2826
10.5821 | -91.2825
10.5824 | -91.2817
10.5828 | -91.2809
10.5831 | -91.2801
10.5834 | -91.2798
10.5838 | -91.2799
10.5841 | -91.2802
10.5844 | -91.2808
10.5847 | -91.2816
10.5851 | -91.2819
10.5854 | -91.2822
10.5857 | -91.2826
10.5861 | -91.2828
10.5864 | -91.2829
10.5867 | -91.283
10.587 | -91.2831
10.5874 | -91.2838
10.5877 | -91.2845
10.588 | -91.2846
10.5884 | -91.2851
10.5887 | -91.2852
10.589 | -91.2865
10.5894 | -91.2881
10.5897 | -91.2883
10.59 | -91.2882
10.5903 | -91.2882
10.5907 | -91.2883
10.591 | -91.2884
10.5913 | -91.2887
10.5917 | -91.2889
10.592 | -91.2888
10.5923 | -91.289
10.5926 | -91.2895
10.593 | -91.2897
10.5933 | -91.2894
10.5936 | -91.2895
10.594 | -91.2896
10.5943 | -91.2901
10.5946 | -91.2899
10.595 | -91.2902
10.5953 | -91.2904
10.5956 | -91.291
10.5959 | -91.2913
10.5963 | -91.2922
10.5966 | -91.2932
10.5969 | -91.2935
10.5973 | -91.2936
10.5976 | -91.2938
10.5979 | -91.294
10.5982 | -91.2947
10.5986 | -91.2951
10.5989 | -91.2954
10.5992 | -91.2956
10.5996 | -91.2953
10.5999 | -91.2953
10.6002 | -91.2961
10.6005 | -91.2966
10.6009 | -91.2971
10.6012 | -91.2964
10.6015 | -91.2952
10.6019 | -91.2953
10.6022 | -91.2958
10.6025 | -91.2964
10.6028 | -91.2972
10.6032 | -91.2975
10.6035 | -91.2975
10.6038 | -91.2974
10.6042 | -91.2973
10.6045 | -91.2975
10.6048 | -91.2974
10.6052 | -91.2976
10.6055 | -91.2987
10.6058 | -91.2999
10.6061 | -91.3005
10.6065 | -91.3009
10.6068 | -91.3012
10.6071 | -91.3022
10.6075 | -91.3027
10.6078 | -91.3038
10.6081 | -91.3038
10.6084 | -91.3036
10.6088 | -91.3029
10.6091 | -91.302
10.6094 | -91.3021
10.6098 | -91.302
10.6101 | -91.3021
10.6104 | -91.3023
10.6107 | -91.3024
10.6111 | -91.3027
10.6114 | -91.3032
10.6117 | -91.3035
10.6121 | -91.3037
10.6124 | -91.3039
10.6127 | -91.3039
10.6131 | -91.3039
10.6134 | -91.3038
10.6137 | -91.304
10.614 | -91.3045
10.6144 | -91.305
10.6147 | -91.3051
10.615 | -91.3053
10.6154 | -91.306
10.6157 | -91.306
10.616 | -91.3063
10.6163 | -91.3064
10.6167 | -91.3068
10.617 | -91.3064
10.6173 | -91.3067
10.6177 | -91.307
10.618 | -91.3072
10.6183 | -91.3073
10.6186 | -91.3075
10.619 | -91.3076
10.6193 | -91.3078
10.6196 | -91.3077
10.62 | -91.308
10.6203 | -91.3084
10.6206 | -91.3085
10.621 | -91.3087
10.6213 | -91.309
10.6216 | -91.309
10.6219 | -91.3093
10.6223 | -91.3099
10.6226 | -91.3103
10.6229 | -91.3104
10.6233 | -91.3105
10.6236 | -91.3107
10.6239 | -91.3109
10.6242 | -91.3111
10.6246 | -91.3112
10.6249 | -91.3115
10.6252 | -91.3119
10.6256 | -91.3125
10.6259 | -91.3126
10.6262 | -91.312
10.6265 | -91.3121
10.6269 | -91.3122
10.6272 | -91.3123
10.6275 | -91.3127
10.6279 | -91.3135
10.6282 | -91.3137
10.6285 | -91.314
10.6289 | -91.3138
10.6292 | -91.313
10.6295 | -91.313
10.6298 | -91.3131
10.6302 | -91.3141
10.6305 | -91.3149
10.6308 | -91.3151
10.6312 | -91.3153
10.6315 | -91.3158
10.6318 | -91.3163
10.6321 | -91.3169
10.6325 | -91.3171
10.6328 | -91.3168
10.6331 | -91.3169
10.6335 | -91.3175
10.6338 | -91.3185
10.6341 | -91.3192
10.6344 | -91.3191
10.6348 | -91.3194
10.6351 | -91.3196
10.6354 | -91.3196
10.6358 | -91.3197
10.6361 | -91.32
10.6364 | -91.3204
10.6367 | -91.3199
10.6371 | -91.3199
10.6374 | -91.3202
10.6377 | -91.3202
10.6381 | -91.3205
10.6384 | -91.3206
10.6387 | -91.3212
10.6391 | -91.3213
10.6394 | -91.321
10.6397 | -91.321
10.64 | -91.3213
10.6404 | -91.3214
10.6407 | -91.3214
10.641 | -91.3215
10.6414 | -91.3217
10.6417 | -91.3217
10.642 | -91.322
10.6423 | -91.3225
10.6427 | -91.3233
10.643 | -91.3236
10.6433 | -91.3238
10.6437 | -91.3238
10.644 | -91.324
10.6443 | -91.3244
10.6447 | -91.3248
10.645 | -91.325
10.6453 | -91.3255
10.6456 | -91.3252
10.646 | -91.3249
10.6463 | -91.3255
10.6466 | -91.3258
10.647 | -91.3259
10.6473 | -91.3261
10.6476 | -91.3264
10.6479 | -91.3265
10.6483 | -91.326
10.6486 | -91.3259
10.6489 | -91.3262
10.6493 | -91.3264
10.6496 | -91.3266
10.6499 | -91.3268
10.6502 | -91.327
10.6506 | -91.3279
10.6509 | -91.3282
10.6512 | -91.3287
10.6516 | -91.3292
10.6519 | -91.3295
10.6522 | -91.33
10.6525 | -91.3305
10.6529 | -91.3309
10.6532 | -91.3312
10.6535 | -91.3313
10.6539 | -91.3316
10.6542 | -91.332
10.6545 | -91.3335
10.6549 | -91.3343
10.6552 | -91.3364
10.6555 | -91.3372
10.6558 | -91.3374
10.6562 | -91.3376
10.6565 | -91.3376
10.6568 | -91.3378
10.6572 | -91.3381
10.6575 | -91.3382
10.6578 | -91.3387
10.6581 | -91.3394
10.6585 | -91.3401
10.6588 | -91.3405
10.6591 | -91.3405
10.6595 | -91.3404
10.6598 | -91.3408
10.6601 | -91.3416
10.6604 | -91.3421
10.6608 | -91.3425
10.6611 | -91.3427
10.6614 | -91.3432
10.6618 | -91.3435
10.6621 | -91.3446
10.6624 | -91.3451
10.6628 | -91.3457
10.6631 | -91.3462
10.6634 | -91.3464
10.6637 | -91.3469
10.6641 | -91.3474
10.6644 | -91.348
10.6647 | -91.3487
10.6651 | -91.349
10.6654 | -91.3491
10.6657 | -91.3493
10.666 | -91.3496
10.6664 | -91.3502
10.6667 | -91.3505
10.667 | -91.3507
10.6674 | -91.3509
10.6677 | -91.3507
10.668 | -91.351
10.6683 | -91.3519
10.6687 | -91.3521
10.669 | -91.352
10.6693 | -91.3519
10.6697 | -91.3522
10.67 | -91.3527
10.6703 | -91.3535
10.6707 | -91.3542
10.671 | -91.3549
10.6713 | -91.3556
10.6716 | -91.356
10.672 | -91.3563
10.6723 | -91.3567
10.6726 | -91.3568
10.673 | -91.3568
10.6733 | -91.3571
10.6736 | -91.3574
10.6739 | -91.3579
10.6743 | -91.3583
10.6746 | -91.3587
10.6749 | -91.3591
10.6753 | -91.3596
10.6756 | -91.3597
10.6759 | -91.3596
10.6762 | -91.3596
10.6766 | -91.3597
10.6769 | -91.3599
10.6772 | -91.3607
10.6776 | -91.3618
10.6779 | -91.3624
10.6782 | -91.3627
10.6786 | -91.3633
10.6789 | -91.363
10.6792 | -91.3629
10.6795 | -91.363
10.6799 | -91.3634
10.6802 | -91.3642
10.6805 | -91.3645
10.6809 | -91.3648
10.6812 | -91.3654
10.6815 | -91.3659
10.6818 | -91.3661
10.6822 | -91.366
10.6825 | -91.3657
10.6828 | -91.3658
10.6832 | -91.3661
10.6835 | -91.3671
10.6838 | -91.3676
10.6841 | -91.3682
10.6845 | -91.3685
10.6848 | -91.3688
10.6851 | -91.3694
10.6855 | -91.37
10.6858 | -91.3701
10.6861 | -91.3702
10.6865 | -91.3707
10.6868 | -91.3711
10.6871 | -91.3713
10.6874 | -91.372
10.6878 | -91.3724
10.6881 | -91.3727
10.6884 | -91.3729
10.6888 | -91.373
10.6891 | -91.3731
10.6894 | -91.3735
10.6897 | -91.3738
10.6901 | -91.3744
10.6904 | -91.375
10.6907 | -91.3752
10.6911 | -91.3756
10.6914 | -91.3757
10.6917 | -91.3755
10.692 | -91.3758
10.6924 | -91.3761
10.6927 | -91.3761
10.693 | -91.3763
10.6934 | -91.3768
10.6937 | -91.3775
10.694 | -91.3779
10.6943 | -91.378
10.6947 | -91.3782
10.695 | -91.3785
10.6953 | -91.3789
10.6957 | -91.3794
10.696 | -91.3801
10.6963 | -91.3805
10.6967 | -91.381
10.697 | -91.3813
10.6973 | -91.3822
10.6976 | -91.3833
10.698 | -91.384
10.6983 | -91.3844
10.6986 | -91.3846
10.699 | -91.385
10.6993 | -91.3848
10.6996 | -91.3862
10.6999 | -91.3876
10.7003 | -91.3878
10.7006 | -91.3884
10.7009 | -91.3884
10.7013 | -91.3884
10.7016 | -91.389
10.7019 | -91.3895
10.7022 | -91.3897
10.7026 | -91.3905
10.7029 | -91.3906
10.7032 | -91.3906
10.7036 | -91.3915
10.7039 | -91.3938
10.7042 | -91.395
10.7046 | -91.3955
10.7049 | -91.3959
10.7052 | -91.3964
10.7055 | -91.3968
10.7059 | -91.3973
10.7062 | -91.3975
10.7065 | -91.3976
10.7069 | -91.3981
10.7072 | -91.3985
10.7075 | -91.399
10.7078 | -91.3995
10.7082 | -91.3997
10.7085 | -91.3998
10.7088 | -91.4002
10.7092 | -91.4004
10.7095 | -91.4005
10.7098 | -91.401
10.7101 | -91.4016
10.7105 | -91.402
10.7108 | -91.4023
10.7111 | -91.403
10.7115 | -91.4037
10.7118 | -91.4044
10.7121 | -91.4052
10.7125 | -91.4056
10.7128 | -91.406
10.7131 | -91.4063
10.7134 | -91.4068
10.7138 | -91.4071
10.7141 | -91.4074
10.7144 | -91.4078
10.7148 | -91.4082
10.7151 | -91.4086
10.7154 | -91.4091
10.7157 | -91.4096
10.7161 | -91.4102
10.7164 | -91.4104
10.7167 | -91.4109
10.7171 | -91.4116
10.7174 | -91.4117
10.7177 | -91.4119
10.718 | -91.4121
10.7184 | -91.4127
10.7187 | -91.4134
10.719 | -91.4139
10.7194 | -91.4143
10.7197 | -91.4147
10.72 | -91.4151
10.7204 | -91.4159
10.7207 | -91.4165
10.721 | -91.4169
10.7213 | -91.4179
10.7217 | -91.4185
10.722 | -91.419
10.7223 | -91.4194
10.7227 | -91.4198
10.723 | -91.4204
10.7233 | -91.4215
10.7236 | -91.4226
10.724 | -91.4234
10.7243 | -91.4241
10.7246 | -91.4246
10.725 | -91.4253
10.7253 | -91.4257
10.7256 | -91.4265
10.7259 | -91.4274
10.7263 | -91.4284
10.7266 | -91.4288
10.7269 | -91.429
10.7273 | -91.4293
10.7276 | -91.4299
10.7279 | -91.4303
10.7283 | -91.4298
10.7286 | -91.4299
10.7289 | -91.4305
10.7292 | -91.431
10.7296 | -91.4312
10.7299 | -91.4316
10.7302 | -91.4322
10.7306 | -91.4325
10.7309 | -91.4333
10.7312 | -91.4337
10.7315 | -91.4342
10.7319 | -91.4345
10.7322 | -91.4349
10.7325 | -91.4352
10.7329 | -91.4358
10.7332 | -91.4364
10.7335 | -91.4371
10.7338 | -91.4377
10.7342 | -91.4381
10.7345 | -91.4386
10.7348 | -91.4392
10.7352 | -91.4398
10.7355 | -91.4402
10.7358 | -91.4407
10.7362 | -91.4414
10.7365 | -91.4425
10.7368 | -91.4436
10.7371 | -91.4444
10.7375 | -91.4449
10.7378 | -91.4455
10.7381 | -91.4463
10.7385 | -91.447
10.7388 | -91.4474
10.7391 | -91.448
10.7394 | -91.4486
10.7398 | -91.4491
10.7401 | -91.4492
10.7404 | -91.4497
10.7408 | -91.45
10.7411 | -91.4505
10.7414 | -91.4507
10.7417 | -91.4507
10.7421 | -91.4513
10.7424 | -91.4523
10.7427 | -91.4528
10.7431 | -91.4533
10.7434 | -91.4538
10.7437 | -91.4543
10.744 | -91.4548
10.7444 | -91.4553
10.7447 | -91.4562
10.745 | -91.4569
10.7454 | -91.4573
10.7457 | -91.458
10.746 | -91.4587
10.7464 | -91.4592
10.7467 | -91.4597
10.747 | -91.4602
10.7473 | -91.4604
10.7477 | -91.4609
10.748 | -91.4615
10.7483 | -91.4621
10.7487 | -91.4632
10.749 | -91.464
10.7493 | -91.4642
10.7496 | -91.4647
10.75 | -91.4652
10.7503 | -91.4658
10.7506 | -91.4664
10.751 | -91.4669
10.7513 | -91.4673
10.7516 | -91.4679
10.7519 | -91.4686
10.7523 | -91.4691
10.7526 | -91.4697
10.7529 | -91.4701
10.7533 | -91.4701
10.7536 | -91.4705
10.7539 | -91.4713
10.7543 | -91.472
10.7546 | -91.4727
10.7549 | -91.4741
10.7552 | -91.4754
10.7556 | -91.4759
10.7559 | -91.4766
10.7562 | -91.4771
10.7566 | -91.4769
10.7569 | -91.4774
10.7572 | -91.4785
10.7575 | -91.4791
10.7579 | -91.4797
10.7582 | -91.4803
10.7585 | -91.4814
10.7589 | -91.4832
10.7592 | -91.4843
10.7595 | -91.4859
10.7598 | -91.486
10.7602 | -91.4861
10.7605 | -91.4867
10.7608 | -91.487
10.7612 | -91.4871
10.7615 | -91.4874
10.7618 | -91.4881
10.7622 | -91.4888
10.7625 | -91.4894
10.7628 | -91.4906
10.7631 | -91.4912
10.7635 | -91.4921
10.7638 | -91.4931
10.7641 | -91.4948
10.7645 | -91.4958
10.7648 | -91.4966
10.7651 | -91.4971
10.7654 | -91.4977
10.7658 | -91.4982
10.7661 | -91.4982
10.7664 | -91.499
10.7668 | -91.5007
10.7671 | -91.5012
10.7674 | -91.5017
10.7677 | -91.5022
10.7681 | -91.5026
10.7684 | -91.5026
10.7687 | -91.5028
10.7691 | -91.5031
10.7694 | -91.5042
10.7697 | -91.5044
10.7701 | -91.5044
10.7704 | -91.5048
10.7707 | -91.5053
10.771 | -91.5064
10.7714 | -91.5076
10.7717 | -91.5087
10.772 | -91.5093
10.7724 | -91.5098
10.7727 | -91.5102
10.773 | -91.5104
10.7733 | -91.5105
10.7737 | -91.5113
10.774 | -91.512
10.7743 | -91.5127
10.7747 | -91.5133
10.775 | -91.5139
10.7753 | -91.5144
10.7756 | -91.515
10.776 | -91.5155
10.7763 | -91.5161
10.7766 | -91.5171
10.777 | -91.5179
10.7773 | -91.5184
10.7776 | -91.5192
10.778 | -91.5197
10.7783 | -91.5204
10.7786 | -91.5212
10.7789 | -91.5221
10.7793 | -91.5219
10.7796 | -91.522
10.7799 | -91.5225
10.7803 | -91.523
10.7806 | -91.5237
10.7809 | -91.5242
10.7812 | -91.5256
10.7816 | -91.5265
10.7819 | -91.527
10.7822 | -91.5275
10.7826 | -91.528
10.7829 | -91.5287
10.7832 | -91.5293
10.7835 | -91.5298
10.7839 | -91.5305
10.7842 | -91.531
10.7845 | -91.5314
10.7849 | -91.5318
10.7852 | -91.5323
10.7855 | -91.533
10.7858 | -91.5337
10.7862 | -91.534
10.7865 | -91.5345
10.7868 | -91.5351
10.7872 | -91.5362
10.7875 | -91.5366
10.7878 | -91.5375
10.7882 | -91.5381
10.7885 | -91.5384
10.7888 | -91.5384
10.7891 | -91.5386
10.7895 | -91.5391
10.7898 | -91.5398
10.7901 | -91.5405
10.7905 | -91.5412
10.7908 | -91.5414
10.7911 | -91.5416
10.7914 | -91.5425
10.7918 | -91.5436
10.7921 | -91.5441
10.7924 | -91.5444
10.7928 | -91.5445
10.7931 | -91.5451
10.7934 | -91.5453
10.7937 | -91.5456
10.7941 | -91.546
10.7944 | -91.5459
10.7947 | -91.5465
10.7951 | -91.547
10.7954 | -91.5478
10.7957 | -91.5484
10.7961 | -91.5491
10.7964 | -91.5499
10.7967 | -91.5505
10.797 | -91.5511
10.7974 | -91.5518
10.7977 | -91.5522
10.798 | -91.5527
10.7984 | -91.5533
10.7987 | -91.5542
10.799 | -91.5554
10.7993 | -91.5555
10.7997 | -91.5562
10.8 | -91.5569
10.8003 | -91.5571
10.8007 | -91.5569
10.801 | -91.5574
10.8013 | -91.5574
10.8016 | -91.5578
10.802 | -91.5587
10.8023 | -91.5591
10.8026 | -91.559
10.803 | -91.5594
10.8033 | -91.5599
10.8036 | -91.5604
10.804 | -91.5605
10.8043 | -91.5608
10.8046 | -91.5614
10.8049 | -91.562
10.8053 | -91.5626
10.8056 | -91.5632
10.8059 | -91.5642
10.8063 | -91.565
10.8066 | -91.5659
10.8069 | -91.5669
10.8072 | -91.5675
10.8076 | -91.568
10.8079 | -91.5686
10.8082 | -91.5694
10.8086 | -91.5702
10.8089 | -91.5706
10.8092 | -91.5715
10.8095 | -91.5723
10.8099 | -91.5726
10.8102 | -91.5735
10.8105 | -91.5744
10.8109 | -91.5745
10.8112 | -91.5749
10.8115 | -91.5754
10.8119 | -91.576
10.8122 | -91.5768
10.8125 | -91.5775
10.8128 | -91.5778
10.8132 | -91.5784
10.8135 | -91.5789
10.8138 | -91.5793
10.8142 | -91.58
10.8145 | -91.5809
10.8148 | -91.5811
10.8151 | -91.5812
10.8155 | -91.5819
10.8158 | -91.5828
10.8161 | -91.5834
10.8165 | -91.5837
10.8168 | -91.5842
10.8171 | -91.5848
10.8174 | -91.5856
10.8178 | -91.5866
10.8181 | -91.5871
10.8184 | -91.5874
10.8188 | -91.5878
10.8191 | -91.5884
10.8194 | -91.589
10.8198 | -91.5895
10.8201 | -91.5902
10.8204 | -91.5905
10.8207 | -91.5904
10.8211 | -91.5908
10.8214 | -91.5913
10.8217 | -91.5917
10.8221 | -91.5919
10.8224 | -91.5922
10.8227 | -91.5928
10.823 | -91.5934
10.8234 | -91.594
10.8237 | -91.5949
10.824 | -91.5955
10.8244 | -91.5963
10.8247 | -91.5976
10.825 | -91.5982
10.8253 | -91.5987
10.8257 | -91.599
10.826 | -91.5988
10.8263 | -91.5988
10.8267 | -91.5993
10.827 | -91.5995
10.8273 | -91.5992
10.8277 | -91.5996
10.828 | -91.6002
10.8283 | -91.6003
10.8286 | -91.6011
10.829 | -91.6018
10.8293 | -91.6018
10.8296 | -91.6022
10.83 | -91.6027
10.8303 | -91.6032
10.8306 | -91.6035
10.8309 | -91.604
10.8313 | -91.6048
10.8316 | -91.6055
10.8319 | -91.6058
10.8323 | -91.6067
10.8326 | -91.6072
10.8329 | -91.6077
10.8332 | -91.6081
10.8336 | -91.6086
10.8339 | -91.6091
10.8342 | -91.6096
10.8346 | -91.6101
10.8349 | -91.6108
10.8352 | -91.6116
10.8355 | -91.6124
10.8359 | -91.6129
10.8362 | -91.6134
10.8365 | -91.6134
10.8369 | -91.6136
10.8372 | -91.6143
10.8375 | -91.6148
10.8379 | -91.6155
10.8382 | -91.6162
10.8385 | -91.6168
10.8388 | -91.6171
10.8392 | -91.6177
10.8395 | -91.6175
10.8398 | -91.6171
10.8402 | -91.6172
10.8405 | -91.6176
10.8408 | -91.6176
10.8411 | -91.6179
10.8415 | -91.6183
10.8418 | -91.6187
10.8421 | -91.6194
10.8425 | -91.6196
10.8428 | -91.6194
10.8431 | -91.6196
10.8434 | -91.6207
10.8438 | -91.6218
10.8441 | -91.6223
10.8444 | -91.6228
10.8448 | -91.6228
10.8451 | -91.6233
10.8454 | -91.6235
10.8458 | -91.624
10.8461 | -91.6248
10.8464 | -91.6254
10.8467 | -91.6258
10.8471 | -91.6264
10.8474 | -91.6266
10.8477 | -91.627
10.8481 | -91.6276
10.8484 | -91.6282
10.8487 | -91.6288
10.849 | -91.6295
10.8494 | -91.6304
10.8497 | -91.6311
10.85 | -91.6319
10.8504 | -91.6322
10.8507 | -91.6325
10.851 | -91.6334
10.8513 | -91.6338
10.8517 | -91.6341
10.852 | -91.6345
10.8523 | -91.6351
10.8527 | -91.6354
10.853 | -91.6357
10.8533 | -91.6369
10.8537 | -91.6378
10.854 | -91.6383
10.8543 | -91.6389
10.8546 | -91.6395
10.855 | -91.6397
10.8553 | -91.6401
10.8556 | -91.6403
10.856 | -91.6404
10.8563 | -91.6411
10.8566 | -91.6414
10.8569 | -91.6417
10.8573 | -91.6422
10.8576 | -91.643
10.8579 | -91.6435
10.8583 | -91.6439
10.8586 | -91.6446
10.8589 | -91.6453
10.8592 | -91.6457
10.8596 | -91.646
10.8599 | -91.6465
10.8602 | -91.6472
10.8606 | -91.6483
10.8609 | -91.6481
10.8612 | -91.6483
10.8616 | -91.6491
10.8619 | -91.6495
10.8622 | -91.65
10.8625 | -91.6505
10.8629 | -91.6502
10.8632 | -91.6499
10.8635 | -91.65
10.8639 | -91.649
10.8642 | -91.6493
10.8645 | -91.6498
10.8648 | -91.6501
10.8652 | -91.6507
10.8655 | -91.6511
10.8658 | -91.6511
10.8662 | -91.6515
10.8665 | -91.6517
10.8668 | -91.6513
10.8671 | -91.6516
10.8675 | -91.652
10.8678 | -91.6523
10.8681 | -91.6526
10.8685 | -91.6531
10.8688 | -91.6535
10.8691 | -91.6539
10.8695 | -91.6543
10.8698 | -91.6548
10.8701 | -91.6555
10.8704 | -91.6559
10.8708 | -91.6563
10.8711 | -91.6569
10.8714 | -91.6576
10.8718 | -91.6579
10.8721 | -91.6583
10.8724 | -91.6592
10.8727 | -91.66
10.8731 | -91.6603
10.8734 | -91.6609
10.8737 | -91.6612
10.8741 | -91.6618
10.8744 | -91.6623
10.8747 | -91.6621
10.875 | -91.6621
10.8754 | -91.6623
10.8757 | -91.6627
10.876 | -91.6632
10.8764 | -91.6632
10.8767 | -91.6636
10.877 | -91.6642
10.8773 | -91.6651
10.8777 | -91.6657
10.878 | -91.6661
10.8783 | -91.666
10.8787 | -91.6654
10.879 | -91.6649
10.8793 | -91.6651
10.8797 | -91.6658
10.88 | -91.6663
10.8803 | -91.6668
10.8806 | -91.6671
10.881 | -91.6675
10.8813 | -91.6681
10.8816 | -91.6689
10.882 | -91.6692
10.8823 | -91.6699
10.8826 | -91.6703
10.8829 | -91.6709
10.8833 | -91.6712
10.8836 | -91.6715
10.8839 | -91.6723
10.8843 | -91.6729
10.8846 | -91.6732
10.8849 | -91.6736
10.8852 | -91.674
10.8856 | -91.6744
10.8859 | -91.6748
10.8862 | -91.6753
10.8866 | -91.6756
10.8869 | -91.6761
10.8872 | -91.6764
10.8876 | -91.6765
10.8879 | -91.6769
10.8882 | -91.6772
10.8885 | -91.6776
10.8889 | -91.678
10.8892 | -91.6786
10.8895 | -91.6792
10.8899 | -91.6799
10.8902 | -91.6814
10.8905 | -91.6818
10.8908 | -91.6826
10.8912 | -91.6837
10.8915 | -91.6841
10.8918 | -91.6844
10.8922 | -91.6849
10.8925 | -91.6852
10.8928 | -91.6853
10.8931 | -91.6858
10.8935 | -91.6865
10.8938 | -91.6867
10.8941 | -91.6869
10.8945 | -91.6873
10.8948 | -91.6879
10.8951 | -91.6883
10.8955 | -91.6888
10.8958 | -91.6893
10.8961 | -91.6897
10.8964 | -91.6901
10.8968 | -91.6905
10.8971 | -91.691
10.8974 | -91.6916
10.8978 | -91.6918
10.8981 | -91.6921
10.8984 | -91.6931
10.8987 | -91.6937
10.8991 | -91.6939
10.8994 | -91.6943
10.8997 | -91.6952
10.9001 | -91.6967
10.9004 | -91.6976
10.9007 | -91.698
10.901 | -91.6985
10.9014 | -91.6991
10.9017 | -91.6993
10.902 | -91.7002
10.9024 | -91.701
10.9027 | -91.7013
10.903 | -91.7016
10.9034 | -91.7018
10.9037 | -91.7017
10.904 | -91.7018
10.9043 | -91.702
10.9047 | -91.7023
10.905 | -91.7028
10.9053 | -91.7035
10.9057 | -91.7043
10.906 | -91.7053
10.9063 | -91.7057
10.9066 | -91.706
10.907 | -91.7063
10.9073 | -91.7057
10.9076 | -91.7055
10.908 | -91.7057
10.9083 | -91.706
10.9086 | -91.7062
10.9089 | -91.707
10.9093 | -91.708
10.9096 | -91.7085
10.9099 | -91.7089
10.9103 | -91.7092
10.9106 | -91.7098
10.9109 | -91.7108
10.9113 | -91.7112
10.9116 | -91.712
10.9119 | -91.7123
10.9122 | -91.7124
10.9126 | -91.7126
10.9129 | -91.713
10.9132 | -91.7134
10.9136 | -91.7137
10.9139 | -91.714
10.9142 | -91.7141
10.9145 | -91.714
10.9149 | -91.7142
10.9152 | -91.7145
10.9155 | -91.7146
10.9159 | -91.715
10.9162 | -91.7159
10.9165 | -91.7167
10.9168 | -91.7165
10.9172 | -91.7166
10.9175 | -91.717
10.9178 | -91.7173
10.9182 | -91.7171
10.9185 | -91.7168
10.9188 | -91.7166
10.9192 | -91.7167
10.9195 | -91.7175
10.9198 | -91.7181
10.9201 | -91.7186
10.9205 | -91.7189
10.9208 | -91.7189
10.9211 | -91.7198
10.9215 | -91.7201
10.9218 | -91.7203
10.9221 | -91.7205
10.9224 | -91.7204
10.9228 | -91.7202
10.9231 | -91.7203
10.9234 | -91.7204
10.9238 | -91.7202
10.9241 | -91.7197
10.9244 | -91.7197
10.9247 | -91.7197
10.9251 | -91.7198
10.9254 | -91.7196
10.9257 | -91.7198
10.9261 | -91.72
10.9264 | -91.7202
10.9267 | -91.7205
10.927 | -91.721
10.9274 | -91.7217
10.9277 | -91.7215
10.928 | -91.7213
10.9284 | -91.721
10.9287 | -91.7214
10.929 | -91.722
10.9294 | -91.7225
10.9297 | -91.7226
10.93 | -91.7225
10.9303 | -91.7223
10.9307 | -91.7225
10.931 | -91.7226
10.9313 | -91.7225
10.9317 | -91.7224
10.932 | -91.7219
10.9323 | -91.7213
10.9326 | -91.7215
10.933 | -91.7223
10.9333 | -91.7229
10.9336 | -91.7232
10.934 | -91.7236
10.9343 | -91.7235
10.9346 | -91.7231
10.9349 | -91.7229
10.9353 | -91.7227
10.9356 | -91.7226
10.9359 | -91.7225
10.9363 | -91.7226
10.9366 | -91.7223
10.9369 | -91.7217
10.9373 | -91.7219
10.9376 | -91.7219
10.9379 | -91.722
10.9382 | -91.722
10.9386 | -91.7218
10.9389 | -91.7219
10.9392 | -91.722
10.9396 | -91.722
10.9399 | -91.7219
10.9402 | -91.7219
10.9405 | -91.7224
10.9409 | -91.723
10.9412 | -91.723
10.9415 | -91.7235
10.9419 | -91.7236
10.9422 | -91.7237
10.9425 | -91.7237
10.9428 | -91.7236
10.9432 | -91.724
10.9435 | -91.7245
10.9438 | -91.7249
10.9442 | -91.7252
10.9445 | -91.7258
10.9448 | -91.7261
10.9452 | -91.7264
10.9455 | -91.7267
10.9458 | -91.7267
10.9461 | -91.7268
10.9465 | -91.7268
10.9468 | -91.7269
10.9471 | -91.7272
10.9475 | -91.7268
10.9478 | -91.7261
10.9481 | -91.7264
10.9484 | -91.7264
10.9488 | -91.7263
10.9491 | -91.7264
10.9494 | -91.7266
10.9498 | -91.7269
10.9501 | -91.7273
10.9504 | -91.7275
10.9507 | -91.7275
10.9511 | -91.7281
10.9514 | -91.7288
10.9517 | -91.7293
10.9521 | -91.7292
10.9524 | -91.7291
10.9527 | -91.7291
10.9531 | -91.7289
10.9534 | -91.7289
10.9537 | -91.7289
10.954 | -91.7284
10.9544 | -91.728
10.9547 | -91.7285
10.955 | -91.729
10.9554 | -91.7289
10.9557 | -91.7289
10.956 | -91.729
10.9563 | -91.7291
10.9567 | -91.7297
10.957 | -91.73
10.9573 | -91.7298
10.9577 | -91.7298
10.958 | -91.7298
10.9583 | -91.7301
10.9586 | -91.7301
10.959 | -91.7308
10.9593 | -91.7318
10.9596 | -91.732
10.96 | -91.7321
10.9603 | -91.7322
10.9606 | -91.7319
10.961 | -91.7315
10.9613 | -91.7311
10.9616 | -91.7316
10.9619 | -91.7329
10.9623 | -91.7333
10.9626 | -91.7333
10.9629 | -91.7334
10.9633 | -91.7335
10.9636 | -91.734
10.9639 | -91.7342
10.9642 | -91.7344
10.9646 | -91.7344
10.9649 | -91.7344
10.9652 | -91.7344
10.9656 | -91.7345
10.9659 | -91.7341
10.9662 | -91.7339
10.9665 | -91.7341
10.9669 | -91.7345
10.9672 | -91.7347
10.9675 | -91.7349
10.9679 | -91.7349
10.9682 | -91.7347
10.9685 | -91.7348
10.9688 | -91.7347
10.9692 | -91.7346
10.9695 | -91.7346
10.9698 | -91.7348
10.9702 | -91.7352
10.9705 | -91.7352
10.9708 | -91.7349
10.9712 | -91.7349
10.9715 | -91.7349
10.9718 | -91.7352
10.9721 | -91.7352
10.9725 | -91.7351
10.9728 | -91.7349
10.9731 | -91.7345
10.9735 | -91.734
10.9738 | -91.7345
10.9741 | -91.735
10.9744 | -91.7351
10.9748 | -91.7352
10.9751 | -91.7355
10.9754 | -91.7358
10.9758 | -91.7359
10.9761 | -91.7354
10.9764 | -91.7359
10.9767 | -91.7366
10.9771 | -91.7366
10.9774 | -91.7365
10.9777 | -91.7366
10.9781 | -91.7369
10.9784 | -91.7369
10.9787 | -91.7368
10.9791 | -91.7367
10.9794 | -91.737
10.9797 | -91.7374
10.98 | -91.7375
10.9804 | -91.7377
10.9807 | -91.7378
10.981 | -91.7378
10.9814 | -91.7378
10.9817 | -91.737
10.982 | -91.7361
10.9823 | -91.7361
10.9827 | -91.7363
10.983 | -91.7366
10.9833 | -91.7368
10.9837 | -91.7367
10.984 | -91.7364
10.9843 | -91.7363
10.9846 | -91.7363
10.985 | -91.737
10.9853 | -91.7372
10.9856 | -91.7369
10.986 | -91.7366
10.9863 | -91.7362
10.9866 | -91.736
10.987 | -91.7358
10.9873 | -91.7352
10.9876 | -91.7353
10.9879 | -91.7352
10.9883 | -91.7352
10.9886 | -91.7352
10.9889 | -91.7354
10.9893 | -91.7354
10.9896 | -91.7354
10.9899 | -91.7356
10.9902 | -91.7356
10.9906 | -91.7355
10.9909 | -91.7354
10.9912 | -91.7355
10.9916 | -91.7355
10.9919 | -91.7356
10.9922 | -91.7358
10.9925 | -91.7354
10.9929 | -91.7351
10.9932 | -91.7349
10.9935 | -91.7351
10.9939 | -91.7353
10.9942 | -91.7351
10.9945 | -91.7354
10.9949 | -91.7357
10.9952 | -91.7358
10.9955 | -91.7355
10.9958 | -91.7352
10.9962 | -91.7349
10.9965 | -91.7346
10.9968 | -91.7342
10.9972 | -91.7341
10.9975 | -91.734
10.9978 | -91.734
10.9981 | -91.7338
10.9985 | -91.7331
10.9988 | -91.732
10.9991 | -91.7316
10.9995 | -91.7313
10.9998 | -91.7312
11.0001 | -91.7318
11.0004 | -91.7326
11.0008 | -91.7325
11.0011 | -91.7329
11.0014 | -91.7341
11.0018 | -91.7344
11.0021 | -91.734
11.0024 | -91.7336
11.0028 | -91.7337
11.0031 | -91.7333
11.0034 | -91.7334
11.0037 | -91.734
11.0041 | -91.7338
11.0044 | -91.7335
11.0047 | -91.7342
11.0051 | -91.7357
11.0054 | -91.7359
11.0057 | -91.7356
11.006 | -91.7351
11.0064 | -91.7348
11.0067 | -91.7348
11.007 | -91.7349
11.0074 | -91.7345
11.0077 | -91.7341
11.008 | -91.734
11.0083 | -91.7342
11.0087 | -91.7346
11.009 | -91.7344
11.0093 | -91.735
11.0097 | -91.735
11.01 | -91.7344
11.0103 | -91.7333
11.0107 | -91.7327
11.011 | -91.7324
11.0113 | -91.7322
11.0116 | -91.7321
11.012 | -91.7319
11.0123 | -91.7321
11.0126 | -91.7317
11.013 | -91.7314
11.0133 | -91.7309
11.0136 | -91.7303
11.0139 | -91.7296
11.0143 | -91.7293
11.0146 | -91.7293
11.0149 | -91.7292
11.0153 | -91.7289
11.0156 | -91.7288
11.0159 | -91.7288
11.0162 | -91.7287
11.0166 | -91.7284
11.0169 | -91.7282
11.0172 | -91.7279
11.0176 | -91.7271
11.0179 | -91.7269
11.0182 | -91.7267
11.0185 | -91.7263
11.0189 | -91.7257
11.0192 | -91.7258
11.0195 | -91.7255
11.0199 | -91.7253
11.0202 | -91.725
11.0205 | -91.724
11.0209 | -91.7234
11.0212 | -91.7233
11.0215 | -91.7234
11.0218 | -91.7233
11.0222 | -91.7231
11.0225 | -91.723
11.0228 | -91.7228
11.0232 | -91.7224
11.0235 | -91.722
11.0238 | -91.7217
11.0241 | -91.7216
11.0245 | -91.7212
11.0248 | -91.7208
11.0251 | -91.7204
11.0255 | -91.7202
11.0258 | -91.7198
11.0261 | -91.7195
11.0265 | -91.7192
11.0268 | -91.719
11.0271 | -91.7188
11.0274 | -91.7186
11.0278 | -91.7185
11.0281 | -91.7183
11.0284 | -91.7182
11.0288 | -91.718
11.0291 | -91.7179
11.0294 | -91.7177
11.0297 | -91.7178
11.0301 | -91.7187
11.0304 | -91.719
11.0307 | -91.7188
11.0311 | -91.7186
11.0314 | -91.7187
11.0317 | -91.7186
11.032 | -91.7185
11.0324 | -91.7176
11.0327 | -91.7167
11.033 | -91.7161
11.0334 | -91.7159
11.0337 | -91.7158
11.034 | -91.7157
11.0343 | -91.7154
11.0347 | -91.715
11.035 | -91.7145
11.0353 | -91.7142
11.0357 | -91.7146
11.036 | -91.7143
11.0363 | -91.7132
11.0367 | -91.7132
11.037 | -91.7134
11.0373 | -91.7132
11.0376 | -91.7127
11.038 | -91.712
11.0383 | -91.7118
11.0386 | -91.7114
11.039 | -91.7108
11.0393 | -91.7104
11.0396 | -91.7101
11.0399 | -91.7105
11.0403 | -91.7107
11.0406 | -91.7102
11.0409 | -91.7094
11.0413 | -91.7088
11.0416 | -91.709
11.0419 | -91.7094
11.0422 | -91.7093
11.0426 | -91.7086
11.0429 | -91.7085
11.0432 | -91.7084
11.0436 | -91.7088
11.0439 | -91.7091
11.0442 | -91.7084
11.0446 | -91.7083
11.0449 | -91.7083
11.0452 | -91.7083
11.0455 | -91.708
11.0459 | -91.7077
11.0462 | -91.7075
11.0465 | -91.7073
11.0469 | -91.7072
11.0472 | -91.7069
11.0475 | -91.7065
11.0478 | -91.7064
11.0482 | -91.7062
11.0485 | -91.706
11.0488 | -91.7055
11.0492 | -91.7056
11.0495 | -91.7056
11.0498 | -91.7052
11.0501 | -91.7048
11.0505 | -91.7045
11.0508 | -91.7046
11.0511 | -91.7044
11.0515 | -91.7038
11.0518 | -91.7036
11.0521 | -91.7033
11.0525 | -91.703
11.0528 | -91.7027
11.0531 | -91.7028
11.0534 | -91.7025
11.0538 | -91.7019
11.0541 | -91.702
11.0544 | -91.7018
11.0548 | -91.7016
11.0551 | -91.7016
11.0554 | -91.7022
11.0557 | -91.7027
11.0561 | -91.7023
11.0564 | -91.7019
11.0567 | -91.7018
11.0571 | -91.7014
11.0574 | -91.7012
11.0577 | -91.7019
11.058 | -91.7018
11.0584 | -91.7015
11.0587 | -91.7002
11.059 | -91.6991
11.0594 | -91.699
11.0597 | -91.6982
11.06 | -91.6973
11.0604 | -91.6968
11.0607 | -91.6964
11.061 | -91.6961
11.0613 | -91.6958
11.0617 | -91.6956
11.062 | -91.6953
11.0623 | -91.6948
11.0627 | -91.6946
11.063 | -91.6941
11.0633 | -91.6929
11.0636 | -91.6925
11.064 | -91.6922
11.0643 | -91.6917
11.0646 | -91.6912
11.065 | -91.6901
11.0653 | -91.6894
11.0656 | -91.6892
11.0659 | -91.6888
11.0663 | -91.6891
11.0666 | -91.6896
11.0669 | -91.6893
11.0673 | -91.689
11.0676 | -91.6887
11.0679 | -91.6884
11.0683 | -91.6879
11.0686 | -91.6874
11.0689 | -91.6873
11.0692 | -91.6869
11.0696 | -91.6865
11.0699 | -91.6866
11.0702 | -91.687
11.0706 | -91.6866
11.0709 | -91.6864
11.0712 | -91.6864
11.0715 | -91.6861
11.0719 | -91.6858
11.0722 | -91.6852
11.0725 | -91.6849
11.0729 | -91.6842
11.0732 | -91.6831
11.0735 | -91.6822
11.0738 | -91.6815
11.0742 | -91.6812
11.0745 | -91.6814
11.0748 | -91.6814
11.0752 | -91.6808
11.0755 | -91.6804
11.0758 | -91.6805
11.0761 | -91.6803
11.0765 | -91.68
11.0768 | -91.6796
11.0771 | -91.6793
11.0775 | -91.679
11.0778 | -91.6787
11.0781 | -91.6784
11.0785 | -91.678
11.0788 | -91.6779
11.0791 | -91.6775
11.0794 | -91.677
11.0798 | -91.6768
11.0801 | -91.6772
11.0804 | -91.6773
11.0808 | -91.6772
11.0811 | -91.6771
11.0814 | -91.6767
11.0817 | -91.6763
11.0821 | -91.676
11.0824 | -91.6757
11.0827 | -91.6751
11.0831 | -91.6748
11.0834 | -91.6747
11.0837 | -91.6745
11.084 | -91.674
11.0844 | -91.6736
11.0847 | -91.6733
11.085 | -91.6723
11.0854 | -91.6715
11.0857 | -91.6711
11.086 | -91.6708
11.0864 | -91.6704
11.0867 | -91.6703
11.087 | -91.6702
11.0873 | -91.6699
11.0877 | -91.6695
11.088 | -91.6695
11.0883 | -91.6693
11.0887 | -91.6689
11.089 | -91.6686
11.0893 | -91.668
11.0896 | -91.6674
11.09 | -91.6672
11.0903 | -91.6664
11.0906 | -91.6658
11.091 | -91.6659
11.0913 | -91.6657
11.0916 | -91.6643
11.0919 | -91.664
11.0923 | -91.6639
11.0926 | -91.6634
11.0929 | -91.6626
11.0933 | -91.662
11.0936 | -91.6619
11.0939 | -91.6622
11.0943 | -91.6624
11.0946 | -91.6619
11.0949 | -91.6615
11.0952 | -91.6613
11.0956 | -91.6606
11.0959 | -91.6601
11.0962 | -91.66
11.0966 | -91.6596
11.0969 | -91.6592
11.0972 | -91.6586
11.0975 | -91.658
11.0979 | -91.6578
11.0982 | -91.6574
11.0985 | -91.6572
11.0989 | -91.6573
11.0992 | -91.657
11.0995 | -91.6556
11.0998 | -91.6547
11.1002 | -91.654
11.1005 | -91.6533
11.1008 | -91.653
11.1012 | -91.6527
11.1015 | -91.6524
11.1018 | -91.6524
11.1022 | -91.6522
11.1025 | -91.6518
11.1028 | -91.6512
11.1031 | -91.6512
11.1035 | -91.6507
11.1038 | -91.6498
11.1041 | -91.6504
11.1045 | -91.6509
11.1048 | -91.6509
11.1051 | -91.6504
11.1054 | -91.6501
11.1058 | -91.6499
11.1061 | -91.6492
11.1064 | -91.6485
11.1068 | -91.6482
11.1071 | -91.648
11.1074 | -91.6483
11.1077 | -91.6479
11.1081 | -91.6471
11.1084 | -91.6466
11.1087 | -91.646
11.1091 | -91.645
11.1094 | -91.6442
11.1097 | -91.644
11.1101 | -91.6438
11.1104 | -91.6434
11.1107 | -91.6431
11.111 | -91.6431
11.1114 | -91.643
11.1117 | -91.6426
11.112 | -91.6424
11.1124 | -91.642
11.1127 | -91.6418
11.113 | -91.6414
11.1133 | -91.6411
11.1137 | -91.6411
11.114 | -91.6408
11.1143 | -91.6405
11.1147 | -91.64
11.115 | -91.6396
11.1153 | -91.6391
11.1156 | -91.639
11.116 | -91.6387
11.1163 | -91.6383
11.1166 | -91.6377
11.117 | -91.6372
11.1173 | -91.6363
11.1176 | -91.6354
11.118 | -91.6346
11.1183 | -91.6339
11.1186 | -91.633
11.1189 | -91.6325
11.1193 | -91.6323
11.1196 | -91.6315
11.1199 | -91.6312
11.1203 | -91.6307
11.1206 | -91.6304
11.1209 | -91.6299
11.1212 | -91.6297
11.1216 | -91.6291
11.1219 | -91.6283
11.1222 | -91.6275
11.1226 | -91.627
11.1229 | -91.6266
11.1232 | -91.6262
11.1235 | -91.6257
11.1239 | -91.625
11.1242 | -91.6247
11.1245 | -91.6249
11.1249 | -91.6249
11.1252 | -91.6247
11.1255 | -91.625
11.1258 | -91.6249
11.1262 | -91.6242
11.1265 | -91.6237
11.1268 | -91.6233
11.1272 | -91.6228
11.1275 | -91.6223
11.1278 | -91.6217
11.1282 | -91.6209
11.1285 | -91.6204
11.1288 | -91.6196
11.1291 | -91.6186
11.1295 | -91.6183
11.1298 | -91.6178
11.1301 | -91.6174
11.1305 | -91.6171
11.1308 | -91.6171
11.1311 | -91.6167
11.1314 | -91.6158
11.1318 | -91.6151
11.1321 | -91.6147
11.1324 | -91.6149
11.1328 | -91.6157
11.1331 | -91.6155
11.1334 | -91.6154
11.1337 | -91.6153
11.1341 | -91.6148
11.1344 | -91.6145
11.1347 | -91.6142
11.1351 | -91.6137
11.1354 | -91.6128
11.1357 | -91.6119
11.1361 | -91.611
11.1364 | -91.6107
11.1367 | -91.6105
11.137 | -91.6105
11.1374 | -91.6101
11.1377 | -91.6094
11.138 | -91.609
11.1384 | -91.6086
11.1387 | -91.6081
11.139 | -91.6073
11.1393 | -91.607
11.1397 | -91.6068
11.14 | -91.6064
11.1403 | -91.6059
11.1407 | -91.6054
11.141 | -91.605
11.1413 | -91.6046
11.1416 | -91.6042
11.142 | -91.6037
11.1423 | -91.6032
11.1426 | -91.6028
11.143 | -91.6024
11.1433 | -91.602
11.1436 | -91.6016
11.144 | -91.6012
11.1443 | -91.6007
11.1446 | -91.6002
11.1449 | -91.5997
11.1453 | -91.5992
11.1456 | -91.5989
11.1459 | -91.5983
11.1463 | -91.5978
11.1466 | -91.5973
11.1469 | -91.597
11.1472 | -91.5969
11.1476 | -91.5966
11.1479 | -91.5964
11.1482 | -91.5961
11.1486 | -91.5958
11.1489 | -91.5952
11.1492 | -91.5948
11.1495 | -91.5948
11.1499 | -91.5942
11.1502 | -91.5935
11.1505 | -91.593
11.1509 | -91.593
11.1512 | -91.5924
11.1515 | -91.5916
11.1519 | -91.5912
11.1522 | -91.5912
11.1525 | -91.591
11.1528 | -91.5907
11.1532 | -91.5902
11.1535 | -91.589
11.1538 | -91.5887
11.1542 | -91.5894
11.1545 | -91.5898
11.1548 | -91.5895
11.1551 | -91.5892
11.1555 | -91.5888
11.1558 | -91.5884
11.1561 | -91.588
11.1565 | -91.5876
11.1568 | -91.5872
11.1571 | -91.5869
11.1574 | -91.5867
11.1578 | -91.5863
11.1581 | -91.586
11.1584 | -91.5859
11.1588 | -91.5856
11.1591 | -91.5851
11.1594 | -91.5844
11.1598 | -91.5836
11.1601 | -91.5827
11.1604 | -91.5817
11.1607 | -91.5814
11.1611 | -91.5812
11.1614 | -91.5811
11.1617 | -91.5809
11.1621 | -91.5803
11.1624 | -91.5797
11.1627 | -91.579
11.163 | -91.5788
11.1634 | -91.5789
11.1637 | -91.5789
11.164 | -91.5783
11.1644 | -91.5775
11.1647 | -91.5769
11.165 | -91.576
11.1653 | -91.5753
11.1657 | -91.5748
11.166 | -91.5745
11.1663 | -91.5739
11.1667 | -91.5731
11.167 | -91.5724
11.1673 | -91.5718
11.1677 | -91.5717
11.168 | -91.5714
11.1683 | -91.5713
11.1686 | -91.5706
11.169 | -91.57
11.1693 | -91.5696
11.1696 | -91.5691
11.17 | -91.5688
11.1703 | -91.5684
11.1706 | -91.5684
11.1709 | -91.5687
11.1713 | -91.5691
11.1716 | -91.5695
11.1719 | -91.5697
11.1723 | -91.5694
11.1726 | -91.5691
11.1729 | -91.5689
11.1732 | -91.5689
11.1736 | -91.5691
11.1739 | -91.5688
11.1742 | -91.5682
11.1746 | -91.5678
11.1749 | -91.5673
11.1752 | -91.5668
11.1755 | -91.5662
11.1759 | -91.5652
11.1762 | -91.5648
11.1765 | -91.5645
11.1769 | -91.5642
11.1772 | -91.5638
11.1775 | -91.5629
11.1779 | -91.5623
11.1782 | -91.5619
11.1785 | -91.5616
11.1788 | -91.5615
11.1792 | -91.5613
11.1795 | -91.5607
11.1798 | -91.5602
11.1802 | -91.5598
11.1805 | -91.56
11.1808 | -91.5604
11.1811 | -91.56
11.1815 | -91.5591
11.1818 | -91.558
11.1821 | -91.5575
11.1825 | -91.5572
11.1828 | -91.5566
11.1831 | -91.5559
11.1834 | -91.555
11.1838 | -91.5541
11.1841 | -91.5537
11.1844 | -91.5535
11.1848 | -91.5536
11.1851 | -91.5534
11.1854 | -91.5529
11.1858 | -91.5527
11.1861 | -91.5524
11.1864 | -91.552
11.1867 | -91.5514
11.1871 | -91.551
11.1874 | -91.5503
11.1877 | -91.5497
11.1881 | -91.5494
11.1884 | -91.549
11.1887 | -91.5487
11.189 | -91.5483
11.1894 | -91.5476
11.1897 | -91.547
11.19 | -91.5466
11.1904 | -91.5464
11.1907 | -91.5457
11.191 | -91.5454
11.1913 | -91.545
11.1917 | -91.5446
11.192 | -91.5441
11.1923 | -91.5437
11.1927 | -91.5433
11.193 | -91.5428
11.1933 | -91.5426
11.1937 | -91.542
11.194 | -91.5417
11.1943 | -91.5409
11.1946 | -91.5403
11.195 | -91.5397
11.1953 | -91.5389
11.1956 | -91.5382
11.196 | -91.5381
11.1963 | -91.5381
11.1966 | -91.5377
11.1969 | -91.537
11.1973 | -91.5364
11.1976 | -91.5359
11.1979 | -91.5354
11.1983 | -91.5348
11.1986 | -91.5341
11.1989 | -91.5335
11.1992 | -91.5332
11.1996 | -91.5329
11.1999 | -91.5327
11.2002 | -91.5326
11.2006 | -91.5323
11.2009 | -91.532
11.2012 | -91.5318
11.2016 | -91.5315
11.2019 | -91.5315
11.2022 | -91.5313
11.2025 | -91.5309
11.2029 | -91.5306
11.2032 | -91.5303
11.2035 | -91.5302
11.2039 | -91.5304
11.2042 | -91.5303
11.2045 | -91.5298
11.2048 | -91.5294
11.2052 | -91.5288
11.2055 | -91.5287
11.2058 | -91.5285
11.2062 | -91.5276
11.2065 | -91.527
11.2068 | -91.5271
11.2071 | -91.5266
11.2075 | -91.5264
11.2078 | -91.5265
11.2081 | -91.526
11.2085 | -91.5258
11.2088 | -91.5255
11.2091 | -91.5253
11.2095 | -91.525
11.2098 | -91.525
11.2101 | -91.5246
11.2104 | -91.5242
11.2108 | -91.5236
11.2111 | -91.5231
11.2114 | -91.5227
11.2118 | -91.5225
11.2121 | -91.5222
11.2124 | -91.522
11.2127 | -91.5219
11.2131 | -91.521
11.2134 | -91.5209
11.2137 | -91.521
11.2141 | -91.5206
11.2144 | -91.5202
11.2147 | -91.5201
11.215 | -91.5199
11.2154 | -91.5197
11.2157 | -91.5194
11.216 | -91.5198
11.2164 | -91.5197
11.2167 | -91.519
11.217 | -91.5187
11.2173 | -91.5184
11.2177 | -91.5183
11.218 | -91.5184
11.2183 | -91.5181
11.2187 | -91.5178
11.219 | -91.5171
11.2193 | -91.5165
11.2197 | -91.5162
11.22 | -91.5163
11.2203 | -91.516
11.2206 | -91.5157
11.221 | -91.5154
11.2213 | -91.5151
11.2216 | -91.5148
11.222 | -91.5147
11.2223 | -91.5147
11.2226 | -91.5148
11.2229 | -91.5144
11.2233 | -91.5133
11.2236 | -91.5129
11.2239 | -91.5131
11.2243 | -91.5133
11.2246 | -91.5135
11.2249 | -91.5134
11.2252 | -91.5132
11.2256 | -91.5131
11.2259 | -91.5126
11.2262 | -91.5125
11.2266 | -91.5121
11.2269 | -91.5116
11.2272 | -91.5112
11.2276 | -91.511
11.2279 | -91.5106
11.2282 | -91.5104
11.2285 | -91.5104
11.2289 | -91.5106
11.2292 | -91.5105
11.2295 | -91.5103
11.2299 | -91.5099
11.2302 | -91.5095
11.2305 | -91.5095
11.2308 | -91.5091
11.2312 | -91.5088
11.2315 | -91.5082
11.2318 | -91.5077
11.2322 | -91.5077
11.2325 | -91.5077
11.2328 | -91.5074
11.2331 | -91.5072
11.2335 | -91.5068
11.2338 | -91.5065
11.2341 | -91.5059
11.2345 | -91.5058
11.2348 | -91.5054
11.2351 | -91.5054
11.2355 | -91.5055
11.2358 | -91.5052
11.2361 | -91.5047
11.2364 | -91.5044
11.2368 | -91.5041
11.2371 | -91.5037
11.2374 | -91.5034
11.2378 | -91.5025
11.2381 | -91.5016
11.2384 | -91.5012
11.2387 | -91.501
11.2391 | -91.5004
11.2394 | -91.4999
11.2397 | -91.4996
11.2401 | -91.4993
11.2404 | -91.4987
11.2407 | -91.4982
11.241 | -91.498
11.2414 | -91.4978
11.2417 | -91.4975
11.242 | -91.4971
11.2424 | -91.4972
11.2427 | -91.4972
11.243 | -91.4969
11.2434 | -91.4967
11.2437 | -91.4967
11.244 | -91.4966
11.2443 | -91.4959
11.2447 | -91.4953
11.245 | -91.4951
11.2453 | -91.4949
11.2457 | -91.4948
11.246 | -91.4944
11.2463 | -91.4935
11.2466 | -91.4931
11.247 | -91.4928
11.2473 | -91.4925
11.2476 | -91.4925
11.248 | -91.4921
11.2483 | -91.4918
11.2486 | -91.4915
11.2489 | -91.4907
11.2493 | -91.4903
11.2496 | -91.4899
11.2499 | -91.4897
11.2503 | -91.4895
11.2506 | -91.4894
11.2509 | -91.4894
11.2513 | -91.489
11.2516 | -91.4888
11.2519 | -91.4888
11.2522 | -91.4884
11.2526 | -91.488
11.2529 | -91.488
11.2532 | -91.4879
11.2536 | -91.4876
11.2539 | -91.4876
11.2542 | -91.4874
11.2545 | -91.4871
11.2549 | -91.4867
11.2552 | -91.4869
11.2555 | -91.4867
11.2559 | -91.4867
11.2562 | -91.4868
11.2565 | -91.4867
11.2568 | -91.4865
11.2572 | -91.4858
11.2575 | -91.4857
11.2578 | -91.4855
11.2582 | -91.4848
11.2585 | -91.4842
11.2588 | -91.4836
11.2591 | -91.4834
11.2595 | -91.4833
11.2598 | -91.4832
11.2601 | -91.483
11.2605 | -91.4828
11.2608 | -91.4829
11.2611 | -91.4831
11.2615 | -91.4831
11.2618 | -91.4829
11.2621 | -91.4826
11.2624 | -91.4823
11.2628 | -91.4826
11.2631 | -91.4824
11.2634 | -91.4823
11.2638 | -91.4821
11.2641 | -91.4818
11.2644 | -91.4815
11.2647 | -91.4811
11.2651 | -91.4804
11.2654 | -91.4801
11.2657 | -91.4798
11.2661 | -91.4791
11.2664 | -91.4795
11.2667 | -91.4801
11.267 | -91.4802
11.2674 | -91.4811
11.2677 | -91.4811
11.268 | -91.4809
11.2684 | -91.4806
11.2687 | -91.4802
11.269 | -91.4795
11.2694 | -91.4793
11.2697 | -91.4795
11.27 | -91.4795
11.2703 | -91.4794
11.2707 | -91.4791
11.271 | -91.4789
11.2713 | -91.4788
11.2717 | -91.4788
11.272 | -91.4793
11.2723 | -91.4792
11.2726 | -91.4795
11.273 | -91.4796
11.2733 | -91.479
11.2736 | -91.479
11.274 | -91.4788
11.2743 | -91.4786
11.2746 | -91.4785
11.2749 | -91.4784
11.2753 | -91.4782
11.2756 | -91.478
11.2759 | -91.4774
11.2763 | -91.4771
11.2766 | -91.4769
11.2769 | -91.4768
11.2773 | -91.477
11.2776 | -91.4768
11.2779 | -91.4774
11.2782 | -91.4774
11.2786 | -91.4774
11.2789 | -91.4771
11.2792 | -91.4767
11.2796 | -91.477
11.2799 | -91.4767
11.2802 | -91.476
11.2805 | -91.4756
11.2809 | -91.4758
11.2812 | -91.4758
11.2815 | -91.4755
11.2819 | -91.4754
11.2822 | -91.4753
11.2825 | -91.4754
11.2828 | -91.4752
11.2832 | -91.4748
11.2835 | -91.4752
11.2838 | -91.4751
11.2842 | -91.4749
11.2845 | -91.4749
11.2848 | -91.4749
11.2852 | -91.4747
11.2855 | -91.4746
11.2858 | -91.4744
11.2861 | -91.4744
11.2865 | -91.4747
11.2868 | -91.4747
11.2871 | -91.4745
11.2875 | -91.4747
11.2878 | -91.4748
11.2881 | -91.4748
11.2884 | -91.4747
11.2888 | -91.4745
11.2891 | -91.4744
11.2894 | -91.4745
11.2898 | -91.4748
11.2901 | -91.4749
11.2904 | -91.4747
11.2907 | -91.4748
11.2911 | -91.4741
11.2914 | -91.4735
11.2917 | -91.4734
11.2921 | -91.4733
11.2924 | -91.474
11.2927 | -91.4742
11.2931 | -91.4739
11.2934 | -91.4739
11.2937 | -91.4738
11.294 | -91.4738
11.2944 | -91.4738
11.2947 | -91.4736
11.295 | -91.4735
11.2954 | -91.4735
11.2957 | -91.4732
11.296 | -91.473
11.2963 | -91.4732
11.2967 | -91.4736
11.297 | -91.4735
11.2973 | -91.4729
11.2977 | -91.4727
11.298 | -91.4722
11.2983 | -91.4718
11.2986 | -91.4717
11.299 | -91.4718
11.2993 | -91.4711
11.2996 | -91.471
11.3 | -91.4712
11.3003 | -91.4715
11.3006 | -91.4721
11.301 | -91.472
11.3013 | -91.4717
11.3016 | -91.4716
11.3019 | -91.4715
11.3023 | -91.4716
11.3026 | -91.4717
11.3029 | -91.4718
11.3033 | -91.4716
11.3036 | -91.4719
11.3039 | -91.4721
11.3042 | -91.4723
11.3046 | -91.4725
11.3049 | -91.4724
11.3052 | -91.4723
11.3056 | -91.4722
11.3059 | -91.4721
11.3062 | -91.472
11.3065 | -91.472
11.3069 | -91.4719
11.3072 | -91.4714
11.3075 | -91.4711
11.3079 | -91.4712
11.3082 | -91.4712
11.3085 | -91.4709
11.3088 | -91.4701
11.3092 | -91.4699
11.3095 | -91.4697
11.3098 | -91.4696
11.3102 | -91.4698
11.3105 | -91.4698
11.3108 | -91.4693
11.3112 | -91.4687
11.3115 | -91.4689
11.3118 | -91.4688
11.3121 | -91.4685
11.3125 | -91.4685
11.3128 | -91.4685
11.3131 | -91.4685
11.3135 | -91.4685
11.3138 | -91.4684
11.3141 | -91.4684
11.3144 | -91.4684
11.3148 | -91.4687
11.3151 | -91.4688
11.3154 | -91.4687
11.3158 | -91.4687
11.3161 | -91.469
11.3164 | -91.4692
11.3167 | -91.4691
11.3171 | -91.4689
11.3174 | -91.4683
11.3177 | -91.4678
11.3181 | -91.4679
11.3184 | -91.4679
11.3187 | -91.4679
11.3191 | -91.4681
11.3194 | -91.4683
11.3197 | -91.4689
11.32 | -91.469
11.3204 | -91.469
11.3207 | -91.4689
11.321 | -91.4688
11.3214 | -91.4686
11.3217 | -91.4683
11.322 | -91.4684
11.3223 | -91.4693
11.3227 | -91.4698
11.323 | -91.4697
11.3233 | -91.4696
11.3237 | -91.4688
11.324 | -91.4683
11.3243 | -91.4682
11.3246 | -91.4685
11.325 | -91.4686
11.3253 | -91.4686
11.3256 | -91.4686
11.326 | -91.4686
11.3263 | -91.4686
11.3266 | -91.4685
11.327 | -91.4683
11.3273 | -91.4678
11.3276 | -91.4675
11.3279 | -91.4676
11.3283 | -91.4679
11.3286 | -91.4681
11.3289 | -91.4683
11.3293 | -91.4678
11.3296 | -91.4674
11.3299 | -91.4673
11.3302 | -91.4669
11.3306 | -91.4668
11.3309 | -91.4669
11.3312 | -91.4671
11.3316 | -91.4676
11.3319 | -91.4672
11.3322 | -91.4671
11.3325 | -91.4673
11.3329 | -91.4677
11.3332 | -91.4678
11.3335 | -91.468
11.3339 | -91.468
11.3342 | -91.4683
11.3345 | -91.4683
11.3349 | -91.4682
11.3352 | -91.4683
11.3355 | -91.4685
11.3358 | -91.469
11.3362 | -91.4689
11.3365 | -91.4688
11.3368 | -91.4688
11.3372 | -91.4689
11.3375 | -91.469
11.3378 | -91.4691
11.3381 | -91.4692
11.3385 | -91.4694
11.3388 | -91.4691
11.3391 | -91.4691
11.3395 | -91.4689
11.3398 | -91.4686
11.3401 | -91.4684
11.3404 | -91.4686
11.3408 | -91.4691
11.3411 | -91.4691
11.3414 | -91.469
11.3418 | -91.4691
11.3421 | -91.4689
11.3424 | -91.469
11.3428 | -91.4691
11.3431 | -91.47
11.3434 | -91.471
11.3437 | -91.4718
11.3441 | -91.4723
11.3444 | -91.4722
11.3447 | -91.472
11.3451 | -91.4717
11.3454 | -91.4713
11.3457 | -91.4714
11.346 | -91.4714
11.3464 | -91.4716
11.3467 | -91.4707
11.347 | -91.4699
11.3474 | -91.4699
11.3477 | -91.4697
11.348 | -91.4696
11.3483 | -91.469
11.3487 | -91.4688
11.349 | -91.4689
11.3493 | -91.4689
11.3497 | -91.4687
11.35 | -91.4689
11.3503 | -91.469
11.3506 | -91.4691
11.351 | -91.4688
11.3513 | -91.4688
11.3516 | -91.4688
11.352 | -91.4687
11.3523 | -91.4688
11.3526 | -91.4689
11.353 | -91.4688
11.3533 | -91.4691
11.3536 | -91.4692
11.3539 | -91.4693
11.3543 | -91.4693
11.3546 | -91.47
11.3549 | -91.4702
11.3553 | -91.4703
11.3556 | -91.4703
11.3559 | -91.4704
11.3562 | -91.4707
11.3566 | -91.4707
11.3569 | -91.4709
11.3572 | -91.4712
11.3576 | -91.4718
11.3579 | -91.4715
11.3582 | -91.4707
11.3585 | -91.4706
11.3589 | -91.4709
11.3592 | -91.4709
11.3595 | -91.471
11.3599 | -91.471
11.3602 | -91.4707
11.3605 | -91.4706
11.3609 | -91.4707
11.3612 | -91.4709
11.3615 | -91.4716
11.3618 | -91.4724
11.3622 | -91.4728
11.3625 | -91.473
11.3628 | -91.4732
11.3632 | -91.4741
11.3635 | -91.4756
11.3638 | -91.476
11.3641 | -91.476
11.3645 | -91.4759
11.3648 | -91.4762
11.3651 | -91.4763
11.3655 | -91.4765
11.3658 | -91.4769
11.3661 | -91.4771
11.3664 | -91.477
11.3668 | -91.4767
11.3671 | -91.4765
11.3674 | -91.4761
11.3678 | -91.4759
11.3681 | -91.4759
11.3684 | -91.476
11.3688 | -91.4761
11.3691 | -91.4758
11.3694 | -91.4757
11.3697 | -91.4759
11.3701 | -91.4758
11.3704 | -91.4758
11.3707 | -91.4767
11.3711 | -91.4771
11.3714 | -91.4781
11.3717 | -91.479
11.372 | -91.479
11.3724 | -91.4793
11.3727 | -91.4797
11.373 | -91.4796
11.3734 | -91.4796
11.3737 | -91.4796
11.374 | -91.4797
11.3743 | -91.4796
11.3747 | -91.4798
11.375 | -91.4799
11.3753 | -91.4801
11.3757 | -91.4805
11.376 | -91.4808
11.3763 | -91.481
11.3767 | -91.4815
11.377 | -91.4816
11.3773 | -91.4817
11.3776 | -91.4818
11.378 | -91.4821
11.3783 | -91.4821
11.3786 | -91.4821
11.379 | -91.4822
11.3793 | -91.4825
11.3796 | -91.4828
11.3799 | -91.4831
11.3803 | -91.4835
11.3806 | -91.4836
11.3809 | -91.4833
11.3813 | -91.4835
11.3816 | -91.4836
11.3819 | -91.4836
11.3822 | -91.4837
11.3826 | -91.484
11.3829 | -91.4841
11.3832 | -91.484
11.3836 | -91.4842
11.3839 | -91.4842
11.3842 | -91.4841
11.3846 | -91.4843
11.3849 | -91.4847
11.3852 | -91.4849
11.3855 | -91.4849
11.3859 | -91.4849
11.3862 | -91.4847
11.3865 | -91.4846
11.3869 | -91.4847
11.3872 | -91.4847
11.3875 | -91.4843
11.3878 | -91.4843
11.3882 | -91.4846
11.3885 | -91.4844
11.3888 | -91.4844
11.3892 | -91.4844
11.3895 | -91.4847
11.3898 | -91.4856
11.3901 | -91.4863
11.3905 | -91.4865
11.3908 | -91.4867
11.3911 | -91.4871
11.3915 | -91.487
11.3918 | -91.4868
11.3921 | -91.4872
11.3925 | -91.4872
11.3928 | -91.4876
11.3931 | -91.4893
11.3934 | -91.4898
11.3938 | -91.4898
11.3941 | -91.4899
11.3944 | -91.49
11.3948 | -91.4901
11.3951 | -91.4903
11.3954 | -91.4909
11.3957 | -91.4907
11.3961 | -91.4905
11.3964 | -91.4907
11.3967 | -91.491
11.3971 | -91.4916
11.3974 | -91.4918
11.3977 | -91.492
11.398 | -91.4923
11.3984 | -91.4925
11.3987 | -91.4925
11.399 | -91.4927
11.3994 | -91.4925
11.3997 | -91.492
11.4 | -91.4924
11.4003 | -91.4928
11.4007 | -91.4929
11.401 | -91.4932
11.4013 | -91.494
11.4017 | -91.494
11.402 | -91.4937
11.4023 | -91.494
11.4027 | -91.494
11.403 | -91.4943
11.4033 | -91.4943
11.4036 | -91.4946
11.404 | -91.4951
11.4043 | -91.4955
11.4046 | -91.4956
11.405 | -91.496
11.4053 | -91.4963
11.4056 | -91.4968
11.4059 | -91.497
11.4063 | -91.4969
11.4066 | -91.4975
11.4069 | -91.4973
11.4073 | -91.4972
11.4076 | -91.4973
11.4079 | -91.4975
11.4083 | -91.4975
11.4086 | -91.4976
11.4089 | -91.4976
11.4092 | -91.4973
11.4096 | -91.4965
11.4099 | -91.496
11.4102 | -91.496
11.4106 | -91.496
11.4109 | -91.4967
11.4112 | -91.4973
11.4115 | -91.4971
11.4119 | -91.4974
11.4122 | -91.4978
11.4125 | -91.4983
11.4129 | -91.4984
11.4132 | -91.4986
11.4135 | -91.4996
11.4138 | -91.5003
11.4142 | -91.5004
11.4145 | -91.5008
11.4148 | -91.5009
11.4152 | -91.5011
11.4155 | -91.5014
11.4158 | -91.5017
11.4161 | -91.5016
11.4165 | -91.5019
11.4168 | -91.502
11.4171 | -91.5021
11.4175 | -91.5019
11.4178 | -91.502
11.4181 | -91.5023
11.4185 | -91.5025
11.4188 | -91.5028
11.4191 | -91.5032
11.4194 | -91.5035
11.4198 | -91.5038
11.4201 | -91.504
11.4204 | -91.5043
11.4208 | -91.5044
11.4211 | -91.5046
11.4214 | -91.505
11.4217 | -91.5052
11.4221 | -91.5052
11.4224 | -91.5056
11.4227 | -91.5059
11.4231 | -91.5061
11.4234 | -91.5062
11.4237 | -91.5065
11.424 | -91.5068
11.4244 | -91.5065
11.4247 | -91.5066
11.425 | -91.5069
11.4254 | -91.5075
11.4257 | -91.508
11.426 | -91.5082
11.4264 | -91.5084
11.4267 | -91.5087
11.427 | -91.5089
11.4273 | -91.5092
11.4277 | -91.5093
11.428 | -91.5096
11.4283 | -91.51
11.4287 | -91.5102
11.429 | -91.5102
11.4293 | -91.5102
11.4296 | -91.5101
11.43 | -91.5102
11.4303 | -91.5105
11.4306 | -91.5105
11.431 | -91.5103
11.4313 | -91.51
11.4316 | -91.5103
11.4319 | -91.5104
11.4323 | -91.5107
11.4326 | -91.5111
11.4329 | -91.5112
11.4333 | -91.5114
11.4336 | -91.5118
11.4339 | -91.512
11.4343 | -91.5122
11.4346 | -91.5128
11.4349 | -91.5131
11.4352 | -91.5135
11.4356 | -91.5138
11.4359 | -91.5143
11.4362 | -91.5148
11.4366 | -91.5151
11.4369 | -91.5151
11.4372 | -91.515
11.4375 | -91.5151
11.4379 | -91.5154
11.4382 | -91.5156
11.4385 | -91.5158
11.4389 | -91.5163
11.4392 | -91.5164
11.4395 | -91.5169
11.4398 | -91.5174
11.4402 | -91.5177
11.4405 | -91.5178
11.4408 | -91.5175
11.4412 | -91.5172
11.4415 | -91.5173
11.4418 | -91.5175
11.4422 | -91.5171
11.4425 | -91.5169
11.4428 | -91.5172
11.4431 | -91.5177
11.4435 | -91.5184
11.4438 | -91.5187
11.4441 | -91.5191
11.4445 | -91.5192
11.4448 | -91.5187
11.4451 | -91.5186
11.4454 | -91.5182
11.4458 | -91.5185
11.4461 | -91.5189
11.4464 | -91.5202
11.4468 | -91.5208
11.4471 | -91.5213
11.4474 | -91.5217
11.4477 | -91.5219
11.4481 | -91.5218
11.4484 | -91.5219
11.4487 | -91.522
11.4491 | -91.5225
11.4494 | -91.523
11.4497 | -91.5238
11.45 | -91.525
11.4504 | -91.5258
11.4507 | -91.5261
11.451 | -91.5264
11.4514 | -91.5262
11.4517 | -91.5259
11.452 | -91.5262
11.4524 | -91.5262
11.4527 | -91.5266
11.453 | -91.5268
11.4533 | -91.5271
11.4537 | -91.5274
11.454 | -91.5278
11.4543 | -91.5284
11.4547 | -91.5285
11.455 | -91.5289
11.4553 | -91.5293
11.4556 | -91.5296
11.456 | -91.5297
11.4563 | -91.5294
11.4566 | -91.5297
11.457 | -91.5299
11.4573 | -91.5295
11.4576 | -91.529
11.4579 | -91.5292
11.4583 | -91.5295
11.4586 | -91.5299
11.4589 | -91.5303
11.4593 | -91.5306
11.4596 | -91.531
11.4599 | -91.531
11.4603 | -91.5309
11.4606 | -91.5303
11.4609 | -91.5298
11.4612 | -91.5294
11.4616 | -91.5297
11.4619 | -91.5302
11.4622 | -91.5301
11.4626 | -91.5301
11.4629 | -91.5303
11.4632 | -91.5307
11.4635 | -91.5312
11.4639 | -91.5316
11.4642 | -91.5317
11.4645 | -91.5323
11.4649 | -91.5326
11.4652 | -91.5326
11.4655 | -91.5327
11.4658 | -91.5331
11.4662 | -91.5336
11.4665 | -91.5344
11.4668 | -91.5347
11.4672 | -91.5349
11.4675 | -91.5352
11.4678 | -91.5357
11.4682 | -91.5366
11.4685 | -91.5371
11.4688 | -91.5374
11.4691 | -91.5375
11.4695 | -91.5374
11.4698 | -91.538
11.4701 | -91.5385
11.4705 | -91.5387
11.4708 | -91.5388
11.4711 | -91.539
11.4714 | -91.539
11.4718 | -91.5392
11.4721 | -91.5389
11.4724 | -91.5393
11.4728 | -91.5399
11.4731 | -91.5396
11.4734 | -91.5396
11.4737 | -91.5398
11.4741 | -91.5398
11.4744 | -91.5401
11.4747 | -91.5402
11.4751 | -91.5404
11.4754 | -91.5406
11.4757 | -91.5407
11.4761 | -91.5411
11.4764 | -91.5409
11.4767 | -91.5406
11.477 | -91.5409
11.4774 | -91.5411
11.4777 | -91.5414
11.478 | -91.5419
11.4784 | -91.5425
11.4787 | -91.5432
11.479 | -91.5436
11.4793 | -91.5435
11.4797 | -91.5432
11.48 | -91.5431
11.4803 | -91.543
11.4807 | -91.5428
11.481 | -91.5432
11.4813 | -91.5435
11.4816 | -91.5438
11.482 | -91.5441
11.4823 | -91.5443
11.4826 | -91.5447
11.483 | -91.5447
11.4833 | -91.5448
11.4836 | -91.5449
11.484 | -91.5454
11.4843 | -91.5454
11.4846 | -91.5455
11.4849 | -91.5457
11.4853 | -91.546
11.4856 | -91.546
11.4859 | -91.5457
11.4863 | -91.546
11.4866 | -91.5464
11.4869 | -91.5465
11.4872 | -91.5462
11.4876 | -91.5464
11.4879 | -91.547
11.4882 | -91.5471
11.4886 | -91.5476
11.4889 | -91.5483
11.4892 | -91.5485
11.4895 | -91.5489
11.4899 | -91.5489
11.4902 | -91.5491
11.4905 | -91.5493
11.4909 | -91.5496
11.4912 | -91.5498
11.4915 | -91.5497
11.4919 | -91.5499
11.4922 | -91.5501
11.4925 | -91.5503
11.4928 | -91.5506
11.4932 | -91.5508
11.4935 | -91.5513
11.4938 | -91.5517
11.4942 | -91.552
11.4945 | -91.5525
11.4948 | -91.5532
11.4951 | -91.5533
11.4955 | -91.5536
11.4958 | -91.5537
11.4961 | -91.5537
11.4965 | -91.5537
11.4968 | -91.5539
11.4971 | -91.5543
11.4974 | -91.5548
11.4978 | -91.5553
11.4981 | -91.5556
11.4984 | -91.5556
11.4988 | -91.5562
11.4991 | -91.5574
11.4994 | -91.558
11.4998 | -91.5581
11.5001 | -91.5585
11.5004 | -91.5579
11.5007 | -91.5576
11.5011 | -91.5579
11.5014 | -91.5582
11.5017 | -91.5586
11.5021 | -91.5588
11.5024 | -91.5593
11.5027 | -91.5594
11.503 | -91.5598
11.5034 | -91.5603
11.5037 | -91.5602
11.504 | -91.5606
11.5044 | -91.5608
11.5047 | -91.5607
11.505 | -91.5609
11.5053 | -91.5613
11.5057 | -91.5615
11.506 | -91.5618
11.5063 | -91.5623
11.5067 | -91.5626
11.507 | -91.5629
11.5073 | -91.5633
11.5076 | -91.5636
11.508 | -91.5634
11.5083 | -91.563
11.5086 | -91.5629
11.509 | -91.563
11.5093 | -91.5633
11.5096 | -91.5634
11.51 | -91.5635
11.5103 | -91.5633
11.5106 | -91.5636
11.5109 | -91.5638
11.5113 | -91.5638
11.5116 | -91.5638
11.5119 | -91.5641
11.5123 | -91.5648
11.5126 | -91.5651
11.5129 | -91.5656
11.5132 | -91.5659
11.5136 | -91.5664
11.5139 | -91.567
11.5142 | -91.5675
11.5146 | -91.5677
11.5149 | -91.568
11.5152 | -91.5681
11.5155 | -91.5683
11.5159 | -91.5685
11.5162 | -91.5687
11.5165 | -91.569
11.5169 | -91.5693
11.5172 | -91.5695
11.5175 | -91.5695
11.5179 | -91.5696
11.5182 | -91.5697
11.5185 | -91.5695
11.5188 | -91.5697
11.5192 | -91.57
11.5195 | -91.5702
11.5198 | -91.5708
11.5202 | -91.5712
11.5205 | -91.5716
11.5208 | -91.5718
11.5211 | -91.5721
11.5215 | -91.5725
11.5218 | -91.5732
11.5221 | -91.5738
11.5225 | -91.5741
11.5228 | -91.5741
11.5231 | -91.5743
11.5234 | -91.5747
11.5238 | -91.5747
11.5241 | -91.575
11.5244 | -91.5753
11.5248 | -91.5754
11.5251 | -91.5762
11.5254 | -91.5766
11.5258 | -91.5768
11.5261 | -91.5773
11.5264 | -91.5779
11.5267 | -91.5782
11.5271 | -91.5784
11.5274 | -91.5783
11.5277 | -91.5785
11.5281 | -91.5787
11.5284 | -91.5791
11.5287 | -91.5794
11.529 | -91.5797
11.5294 | -91.5802
11.5297 | -91.581
11.53 | -91.5812
11.5304 | -91.5812
11.5307 | -91.5811
11.531 | -91.5814
11.5313 | -91.5814
11.5317 | -91.5812
11.532 | -91.5812
11.5323 | -91.5813
11.5327 | -91.5813
11.533 | -91.5813
11.5333 | -91.5825
11.5337 | -91.5833
11.534 | -91.5835
11.5343 | -91.5841
11.5346 | -91.5849
11.535 | -91.5855
11.5353 | -91.5857
11.5356 | -91.586
11.536 | -91.5861
11.5363 | -91.5864
11.5366 | -91.587
11.5369 | -91.5875
11.5373 | -91.5884
11.5376 | -91.5889
11.5379 | -91.5891
11.5383 | -91.5891
11.5386 | -91.5892
11.5389 | -91.5892
11.5392 | -91.5892
11.5396 | -91.5895
11.5399 | -91.5902
11.5402 | -91.5905
11.5406 | -91.5904
11.5409 | -91.5905
11.5412 | -91.5907
11.5416 | -91.591
11.5419 | -91.5914
11.5422 | -91.592
11.5425 | -91.5928
11.5429 | -91.5929
11.5432 | -91.5929
11.5435 | -91.5933
11.5439 | -91.5939
11.5442 | -91.5944
11.5445 | -91.5947
11.5448 | -91.5941
11.5452 | -91.5938
11.5455 | -91.5941
11.5458 | -91.5942
11.5462 | -91.5945
11.5465 | -91.5947
11.5468 | -91.5949
11.5471 | -91.5953
11.5475 | -91.5955
11.5478 | -91.5961
11.5481 | -91.5965
11.5485 | -91.5968
11.5488 | -91.5971
11.5491 | -91.5975
11.5495 | -91.5974
11.5498 | -91.5974
11.5501 | -91.5975
11.5504 | -91.5976
11.5508 | -91.5977
11.5511 | -91.5979
11.5514 | -91.598
11.5518 | -91.5981
11.5521 | -91.5983
11.5524 | -91.5984
11.5527 | -91.5988
11.5531 | -91.5991
11.5534 | -91.5993
11.5537 | -91.5994
11.5541 | -91.6005
11.5544 | -91.6011
11.5547 | -91.6013
11.555 | -91.6014
11.5554 | -91.6015
11.5557 | -91.6017
11.556 | -91.6019
11.5564 | -91.602
11.5567 | -91.6022
11.557 | -91.6021
11.5573 | -91.6018
11.5577 | -91.6017
11.558 | -91.6018
11.5583 | -91.6022
11.5587 | -91.6025
11.559 | -91.6028
11.5593 | -91.6029
11.5597 | -91.6032
11.56 | -91.6033
11.5603 | -91.6033
11.5606 | -91.6033
11.561 | -91.6034
11.5613 | -91.6035
11.5616 | -91.6037
11.562 | -91.6039
11.5623 | -91.604
11.5626 | -91.6043
11.5629 | -91.6045
11.5633 | -91.6048
11.5636 | -91.6044
11.5639 | -91.6043
11.5643 | -91.6048
11.5646 | -91.6053
11.5649 | -91.6054
11.5652 | -91.6055
11.5656 | -91.6058
11.5659 | -91.6059
11.5662 | -91.6063
11.5666 | -91.6066
11.5669 | -91.607
11.5672 | -91.6067
11.5676 | -91.6063
11.5679 | -91.6065
11.5682 | -91.6069
11.5685 | -91.6072
11.5689 | -91.6075
11.5692 | -91.6084
11.5695 | -91.609
11.5699 | -91.6094
11.5702 | -91.6102
11.5705 | -91.6103
11.5708 | -91.6103
11.5712 | -91.6101
11.5715 | -91.6104
11.5718 | -91.6105
11.5722 | -91.6107
11.5725 | -91.611
11.5728 | -91.6112
11.5731 | -91.6108
11.5735 | -91.6105
11.5738 | -91.6109
11.5741 | -91.6111
11.5745 | -91.6112
11.5748 | -91.6111
11.5751 | -91.6109
11.5755 | -91.6117
11.5758 | -91.6118
11.5761 | -91.612
11.5764 | -91.6119
11.5768 | -91.6125
11.5771 | -91.6133
11.5774 | -91.6136
11.5778 | -91.6146
11.5781 | -91.6151
11.5784 | -91.6152
11.5787 | -91.6156
11.5791 | -91.6164
11.5794 | -91.6165
11.5797 | -91.6167
11.5801 | -91.6169
11.5804 | -91.6174
11.5807 | -91.6177
11.581 | -91.6178
11.5814 | -91.6181
11.5817 | -91.6182
11.582 | -91.6183
11.5824 | -91.618
11.5827 | -91.6178
11.583 | -91.6179
11.5834 | -91.6179
11.5837 | -91.6179
11.584 | -91.618
11.5843 | -91.6183
11.5847 | -91.6181
11.585 | -91.6183
11.5853 | -91.6182
11.5857 | -91.6183
11.586 | -91.6186
11.5863 | -91.6185
11.5866 | -91.6185
11.587 | -91.6187
11.5873 | -91.6187
11.5876 | -91.6187
11.588 | -91.6192
11.5883 | -91.6196
11.5886 | -91.6195
11.5889 | -91.6193
11.5893 | -91.6195
11.5896 | -91.6196
11.5899 | -91.6195
11.5903 | -91.6197
11.5906 | -91.6198
11.5909 | -91.6194
11.5913 | -91.6191
11.5916 | -91.6193
11.5919 | -91.6197
11.5922 | -91.6199
11.5926 | -91.6199
11.5929 | -91.6198
11.5932 | -91.6201
11.5936 | -91.6203
11.5939 | -91.6207
11.5942 | -91.6206
11.5945 | -91.6205
11.5949 | -91.6204
11.5952 | -91.6205
11.5955 | -91.6207
11.5959 | -91.6209
11.5962 | -91.6216
11.5965 | -91.6219
11.5968 | -91.6221
11.5972 | -91.6222
11.5975 | -91.6225
11.5978 | -91.6229
11.5982 | -91.6238
11.5985 | -91.6246
11.5988 | -91.6247
11.5991 | -91.6247
11.5995 | -91.625
11.5998 | -91.6253
11.6001 | -91.6255
11.6005 | -91.6256
11.6008 | -91.6258
11.6011 | -91.6262
11.6015 | -91.6264
11.6018 | -91.6263
11.6021 | -91.6264
11.6024 | -91.6265
11.6028 | -91.6267
11.6031 | -91.6269
11.6034 | -91.6271
11.6038 | -91.6278
11.6041 | -91.6288
11.6044 | -91.6292
11.6047 | -91.6297
11.6051 | -91.6297
11.6054 | -91.6299
11.6057 | -91.6304
11.6061 | -91.6307
11.6064 | -91.631
11.6067 | -91.6309
11.607 | -91.6306
11.6074 | -91.6307
11.6077 | -91.631
11.608 | -91.6312
11.6084 | -91.6313
11.6087 | -91.6315
11.609 | -91.6316
11.6094 | -91.6318
11.6097 | -91.6319
11.61 | -91.6323
11.6103 | -91.6327
11.6107 | -91.6328
11.611 | -91.633
11.6113 | -91.6331
11.6117 | -91.6332
11.612 | -91.633
11.6123 | -91.6333
11.6126 | -91.6338
11.613 | -91.6337
11.6133 | -91.6334
11.6136 | -91.6335
11.614 | -91.6339
11.6143 | -91.6341
11.6146 | -91.6343
11.6149 | -91.6345
11.6153 | -91.6347
11.6156 | -91.6353
11.6159 | -91.636
11.6163 | -91.6356
11.6166 | -91.6354
11.6169 | -91.6351
11.6173 | -91.6353
11.6176 | -91.6357
11.6179 | -91.636
11.6182 | -91.6363
11.6186 | -91.6371
11.6189 | -91.638
11.6192 | -91.6383
11.6196 | -91.6386
11.6199 | -91.6394
11.6202 | -91.6398
11.6205 | -91.6401
11.6209 | -91.6402
11.6212 | -91.6403
11.6215 | -91.6407
11.6219 | -91.6416
11.6222 | -91.6418
11.6225 | -91.6419
11.6228 | -91.6417
11.6232 | -91.6418
11.6235 | -91.642
11.6238 | -91.6418
11.6242 | -91.6422
11.6245 | -91.6425
11.6248 | -91.6427
11.6252 | -91.6431
11.6255 | -91.6432
11.6258 | -91.6432
11.6261 | -91.6434
11.6265 | -91.6432
11.6268 | -91.6434
11.6271 | -91.6438
11.6275 | -91.6438
11.6278 | -91.6437
11.6281 | -91.6438
11.6284 | -91.644
11.6288 | -91.6435
11.6291 | -91.6429
11.6294 | -91.6428
11.6298 | -91.643
11.6301 | -91.6434
11.6304 | -91.6438
11.6307 | -91.6436
11.6311 | -91.6436
11.6314 | -91.644
11.6317 | -91.6441
11.6321 | -91.6443
11.6324 | -91.6442
11.6327 | -91.6446
11.6331 | -91.6443
11.6334 | -91.6443
11.6337 | -91.6445
11.634 | -91.6447
11.6344 | -91.6448
11.6347 | -91.6446
11.635 | -91.6446
11.6354 | -91.6442
11.6357 | -91.6445
11.636 | -91.6449
11.6363 | -91.6451
11.6367 | -91.6448
11.637 | -91.6449
11.6373 | -91.6455
11.6377 | -91.6458
11.638 | -91.6459
11.6383 | -91.6458
11.6386 | -91.6458
11.639 | -91.6459
11.6393 | -91.6461
11.6396 | -91.6457
11.64 | -91.6456
11.6403 | -91.6455
11.6406 | -91.6456
11.641 | -91.6457
11.6413 | -91.6464
11.6416 | -91.6471
11.6419 | -91.6475
11.6423 | -91.6477
11.6426 | -91.6477
11.6429 | -91.6473
11.6433 | -91.6475
11.6436 | -91.648
11.6439 | -91.6487
11.6442 | -91.6491
11.6446 | -91.6497
11.6449 | -91.6501
11.6452 | -91.6498
11.6456 | -91.6496
11.6459 | -91.6495
11.6462 | -91.6488
11.6465 | -91.6486
11.6469 | -91.649
11.6472 | -91.6491
11.6475 | -91.649
11.6479 | -91.6497
11.6482 | -91.6499
11.6485 | -91.6497
11.6488 | -91.6497
11.6492 | -91.65
11.6495 | -91.65
11.6498 | -91.6502
11.6502 | -91.6504
11.6505 | -91.6505
11.6508 | -91.6509
11.6512 | -91.6507
11.6515 | -91.6509
11.6518 | -91.6505
11.6521 | -91.6504
11.6525 | -91.6504
11.6528 | -91.6506
11.6531 | -91.6506
11.6535 | -91.6506
11.6538 | -91.6505
11.6541 | -91.6504
11.6544 | -91.6505
11.6548 | -91.6505
11.6551 | -91.6501
11.6554 | -91.6498
11.6558 | -91.6497
11.6561 | -91.6497
11.6564 | -91.65
11.6567 | -91.6503
11.6571 | -91.6508
11.6574 | -91.6509
11.6577 | -91.651
11.6581 | -91.6511
11.6584 | -91.651
11.6587 | -91.6508
11.6591 | -91.6507
11.6594 | -91.6507
11.6597 | -91.6508
11.66 | -91.6513
11.6604 | -91.6517
11.6607 | -91.6524
11.661 | -91.6525
11.6614 | -91.6526
11.6617 | -91.652
11.662 | -91.6517
11.6623 | -91.6516
11.6627 | -91.6511
11.663 | -91.6508
11.6633 | -91.6508
11.6637 | -91.6508
11.664 | -91.651
11.6643 | -91.6513
11.6646 | -91.6513
11.665 | -91.6514
11.6653 | -91.6517
11.6656 | -91.6518
11.666 | -91.6517
11.6663 | -91.6517
11.6666 | -91.6518
11.667 | -91.652
11.6673 | -91.6521
11.6676 | -91.6522
11.6679 | -91.6518
11.6683 | -91.6515
11.6686 | -91.6511
11.6689 | -91.6509
11.6693 | -91.651
11.6696 | -91.6511
11.6699 | -91.6513
11.6702 | -91.651
11.6706 | -91.6508
11.6709 | -91.651
11.6712 | -91.6509
11.6716 | -91.6507
11.6719 | -91.6505
11.6722 | -91.6506
11.6725 | -91.6503
11.6729 | -91.6504
11.6732 | -91.6506
11.6735 | -91.6504
11.6739 | -91.6502
11.6742 | -91.6502
11.6745 | -91.6503
11.6749 | -91.6502
11.6752 | -91.6497
11.6755 | -91.6489
11.6758 | -91.6489
11.6762 | -91.649
11.6765 | -91.6487
11.6768 | -91.6484
11.6772 | -91.6487
11.6775 | -91.6489
11.6778 | -91.6488
11.6781 | -91.6487
11.6785 | -91.6485
11.6788 | -91.6483
11.6791 | -91.6481
11.6795 | -91.6481
11.6798 | -91.6481
11.6801 | -91.6482
11.6804 | -91.6476
11.6808 | -91.6478
11.6811 | -91.648
11.6814 | -91.6481
11.6818 | -91.6481
11.6821 | -91.6482
11.6824 | -91.6484
11.6828 | -91.6483
11.6831 | -91.6483
11.6834 | -91.6484
11.6837 | -91.6487
11.6841 | -91.6491
11.6844 | -91.6493
11.6847 | -91.6494
11.6851 | -91.649
11.6854 | -91.6492
11.6857 | -91.6493
11.686 | -91.6493
11.6864 | -91.6493
11.6867 | -91.6491
11.687 | -91.6491
11.6874 | -91.6493
11.6877 | -91.6489
11.688 | -91.6484
11.6883 | -91.6481
11.6887 | -91.6481
11.689 | -91.648
11.6893 | -91.6479
11.6897 | -91.6477
11.69 | -91.6477
11.6903 | -91.6478
11.6906 | -91.6477
11.691 | -91.6478
11.6913 | -91.6484
11.6916 | -91.649
11.692 | -91.6492
11.6923 | -91.6493
11.6926 | -91.6495
11.693 | -91.6496
11.6933 | -91.6496
11.6936 | -91.6495
11.6939 | -91.6494
11.6943 | -91.6493
11.6946 | -91.6491
11.6949 | -91.6491
11.6953 | -91.6489
11.6956 | -91.6487
11.6959 | -91.6487
11.6962 | -91.6485
11.6966 | -91.6483
11.6969 | -91.6484
11.6972 | -91.6482
11.6976 | -91.6482
11.6979 | -91.6481
11.6982 | -91.6481
11.6985 | -91.648
11.6989 | -91.6481
11.6992 | -91.6482
11.6995 | -91.6472
11.6999 | -91.6467
11.7002 | -91.6466
11.7005 | -91.6463
11.7009 | -91.6461
11.7012 | -91.6462
11.7015 | -91.6465
11.7018 | -91.646
11.7022 | -91.6455
11.7025 | -91.6455
11.7028 | -91.6455
11.7032 | -91.6454
11.7035 | -91.6454
11.7038 | -91.6454
11.7041 | -91.6454
11.7045 | -91.6458
11.7048 | -91.6461
11.7051 | -91.6469
11.7055 | -91.647
11.7058 | -91.6468
11.7061 | -91.6469
11.7064 | -91.647
11.7068 | -91.6469
11.7071 | -91.6469
11.7074 | -91.6474
11.7078 | -91.6476
11.7081 | -91.6476
11.7084 | -91.6472
11.7088 | -91.6475
11.7091 | -91.6475
11.7094 | -91.6481
11.7097 | -91.6482
11.7101 | -91.6486
11.7104 | -91.6486
11.7107 | -91.6488
11.7111 | -91.649
11.7114 | -91.6487
11.7117 | -91.6487
11.712 | -91.6485
11.7124 | -91.6484
11.7127 | -91.6479
11.713 | -91.6482
11.7134 | -91.6481
11.7137 | -91.6485
11.714 | -91.6484
11.7143 | -91.6483
11.7147 | -91.6483
11.715 | -91.6482
11.7153 | -91.6483
11.7157 | -91.6483
11.716 | -91.6483
11.7163 | -91.6481
11.7167 | -91.6478
11.717 | -91.6479
11.7173 | -91.6481
11.7176 | -91.6487
11.718 | -91.6493
11.7183 | -91.6496
11.7186 | -91.6494
11.719 | -91.649
11.7193 | -91.6486
11.7196 | -91.6484
11.7199 | -91.6485
11.7203 | -91.6491
11.7206 | -91.6496
11.7209 | -91.6497
11.7213 | -91.6498
11.7216 | -91.6502
11.7219 | -91.6502
11.7222 | -91.6501
11.7226 | -91.6505
11.7229 | -91.6513
11.7232 | -91.6512
11.7236 | -91.6512
11.7239 | -91.6513
11.7242 | -91.6516
11.7246 | -91.6517
11.7249 | -91.6518
11.7252 | -91.6515
11.7255 | -91.6514
11.7259 | -91.6515
11.7262 | -91.6517
11.7265 | -91.652
11.7269 | -91.6519
11.7272 | -91.6519
11.7275 | -91.6514
11.7278 | -91.6514
11.7282 | -91.6512
11.7285 | -91.6511
11.7288 | -91.651
11.7292 | -91.6513
11.7295 | -91.6515
11.7298 | -91.6513
11.7301 | -91.6513
11.7305 | -91.6511
11.7308 | -91.6509
11.7311 | -91.6506
11.7315 | -91.6504
11.7318 | -91.6504
11.7321 | -91.6504
11.7325 | -91.6505
11.7328 | -91.6505
11.7331 | -91.6502
11.7334 | -91.6499
11.7338 | -91.6496
11.7341 | -91.6502
11.7344 | -91.6509
11.7348 | -91.6506
11.7351 | -91.6506
11.7354 | -91.6508
11.7357 | -91.6508
11.7361 | -91.6507
11.7364 | -91.6507
11.7367 | -91.6506
11.7371 | -91.6505
11.7374 | -91.6504
11.7377 | -91.6503
11.738 | -91.6506
11.7384 | -91.6506
11.7387 | -91.6506
11.739 | -91.6504
11.7394 | -91.6502
11.7397 | -91.65
11.74 | -91.6501
11.7403 | -91.6501
11.7407 | -91.6502
11.741 | -91.6502
11.7413 | -91.6507
11.7417 | -91.6512
11.742 | -91.6511
11.7423 | -91.6511
11.7427 | -91.6512
11.743 | -91.6509
11.7433 | -91.6506
11.7436 | -91.6507
11.744 | -91.6507
11.7443 | -91.6508
11.7446 | -91.6509
11.745 | -91.651
11.7453 | -91.6511
11.7456 | -91.6509
11.7459 | -91.6503
11.7463 | -91.6503
11.7466 | -91.6502
11.7469 | -91.6504
11.7473 | -91.6497
11.7476 | -91.6497
11.7479 | -91.6497
11.7482 | -91.6497
11.7486 | -91.6499
11.7489 | -91.6497
11.7492 | -91.6502
11.7496 | -91.6509
11.7499 | -91.6511
11.7502 | -91.6509
11.7506 | -91.6506
11.7509 | -91.6507
11.7512 | -91.651
11.7515 | -91.651
11.7519 | -91.651
11.7522 | -91.651
11.7525 | -91.6512
11.7529 | -91.6509
11.7532 | -91.6509
11.7535 | -91.6509
11.7538 | -91.6514
11.7542 | -91.6516
11.7545 | -91.6516
11.7548 | -91.6515
11.7552 | -91.6515
11.7555 | -91.6513
11.7558 | -91.6513
11.7561 | -91.6515
11.7565 | -91.6517
11.7568 | -91.6517
11.7571 | -91.6517
11.7575 | -91.6518
11.7578 | -91.6518
11.7581 | -91.6516
11.7585 | -91.6515
11.7588 | -91.6514
11.7591 | -91.6513
11.7594 | -91.6512
11.7598 | -91.6512
11.7601 | -91.6512
11.7604 | -91.6511
11.7608 | -91.6508
11.7611 | -91.6506
11.7614 | -91.6508
11.7617 | -91.6512
11.7621 | -91.6515
11.7624 | -91.6514
11.7627 | -91.6514
11.7631 | -91.6513
11.7634 | -91.6513
11.7637 | -91.6512
11.764 | -91.651
11.7644 | -91.6506
11.7647 | -91.651
11.765 | -91.6512
11.7654 | -91.6509
11.7657 | -91.6509
11.766 | -91.6504
11.7664 | -91.6506
11.7667 | -91.6505
11.767 | -91.6505
11.7673 | -91.6503
11.7677 | -91.6503
11.768 | -91.6501
11.7683 | -91.6498
11.7687 | -91.6494
11.769 | -91.6493
11.7693 | -91.6492
11.7696 | -91.6492
11.77 | -91.6492
11.7703 | -91.6488
11.7706 | -91.6486
11.771 | -91.6494
11.7713 | -91.6502
11.7716 | -91.6501
11.7719 | -91.6498
11.7723 | -91.6496
11.7726 | -91.6495
11.7729 | -91.6495
11.7733 | -91.6498
11.7736 | -91.6494
11.7739 | -91.6493
11.7743 | -91.6491
11.7746 | -91.6492
11.7749 | -91.6493
11.7752 | -91.6498
11.7756 | -91.65
11.7759 | -91.6504
11.7762 | -91.6509
11.7766 | -91.6509
11.7769 | -91.6508
11.7772 | -91.6511
11.7775 | -91.6508
11.7779 | -91.6508
11.7782 | -91.6509
11.7785 | -91.6507
11.7789 | -91.6504
11.7792 | -91.6504
11.7795 | -91.6503
11.7798 | -91.6502
11.7802 | -91.65
11.7805 | -91.6496
11.7808 | -91.6495
11.7812 | -91.6493
11.7815 | -91.6489
11.7818 | -91.6481
11.7821 | -91.6477
11.7825 | -91.6483
11.7828 | -91.6485
11.7831 | -91.6486
11.7835 | -91.6484
11.7838 | -91.6484
11.7841 | -91.6485
11.7845 | -91.6487
11.7848 | -91.6489
11.7851 | -91.649
11.7854 | -91.649
11.7858 | -91.649
11.7861 | -91.6496
11.7864 | -91.6494
11.7868 | -91.6493
11.7871 | -91.6494
11.7874 | -91.6492
11.7877 | -91.6493
11.7881 | -91.6494
11.7884 | -91.649
11.7887 | -91.6489
11.7891 | -91.6494
11.7894 | -91.6503
11.7897 | -91.6506
11.7901 | -91.6509
11.7904 | -91.6511
11.7907 | -91.6511
11.791 | -91.6509
11.7914 | -91.6511
11.7917 | -91.6507
11.792 | -91.6502
11.7924 | -91.6505
11.7927 | -91.6509
11.793 | -91.651
11.7933 | -91.6508
11.7937 | -91.6504
11.794 | -91.6504
11.7943 | -91.6506
11.7947 | -91.6511
11.795 | -91.6513
11.7953 | -91.6516
11.7956 | -91.6516
11.796 | -91.6516
11.7963 | -91.6514
11.7966 | -91.6516
11.797 | -91.6518
11.7973 | -91.6516
11.7976 | -91.6517
11.7979 | -91.6516
11.7983 | -91.6514
11.7986 | -91.6514
11.7989 | -91.6509
11.7993 | -91.6508
11.7996 | -91.651
11.7999 | -91.6509
11.8003 | -91.651
11.8006 | -91.6509
11.8009 | -91.6508
11.8012 | -91.6505
11.8016 | -91.6508
11.8019 | -91.6507
11.8022 | -91.6506
11.8026 | -91.6506
11.8029 | -91.6501
11.8032 | -91.6494
11.8035 | -91.649
11.8039 | -91.6494
11.8042 | -91.6491
11.8045 | -91.6485
11.8049 | -91.6478
11.8052 | -91.6477
11.8055 | -91.6479
11.8058 | -91.6476
11.8062 | -91.6475
11.8065 | -91.6472
11.8068 | -91.6471
11.8072 | -91.6471
11.8075 | -91.6471
11.8078 | -91.6468
11.8082 | -91.6469
11.8085 | -91.6472
11.8088 | -91.647
11.8091 | -91.6472
11.8095 | -91.6475
11.8098 | -91.6475
11.8101 | -91.6475
11.8105 | -91.6473
11.8108 | -91.6474
11.8111 | -91.6464
11.8114 | -91.6462
11.8118 | -91.6463
11.8121 | -91.6461
11.8124 | -91.6465
11.8128 | -91.6466
11.8131 | -91.6463
11.8134 | -91.6457
11.8137 | -91.6457
11.8141 | -91.6453
11.8144 | -91.645
11.8147 | -91.6446
11.8151 | -91.6439
11.8154 | -91.6438
11.8157 | -91.6438
11.8161 | -91.6439
11.8164 | -91.644
11.8167 | -91.6436
11.817 | -91.6435
11.8174 | -91.6437
11.8177 | -91.6438
11.818 | -91.644
11.8184 | -91.6441
11.8187 | -91.644
11.819 | -91.6439
11.8193 | -91.6439
11.8197 | -91.6438
11.82 | -91.644
11.8203 | -91.6438
11.8207 | -91.6439
11.821 | -91.6439
11.8213 | -91.6438
11.8216 | -91.6438
11.822 | -91.6437
11.8223 | -91.6433
11.8226 | -91.6431
11.823 | -91.6427
11.8233 | -91.6419
11.8236 | -91.6413
11.824 | -91.6408
11.8243 | -91.6406
11.8246 | -91.6404
11.8249 | -91.6405
11.8253 | -91.641
11.8256 | -91.6412
11.8259 | -91.641
11.8263 | -91.6406
11.8266 | -91.6402
11.8269 | -91.6401
11.8272 | -91.6402
11.8276 | -91.6404
11.8279 | -91.6403
11.8282 | -91.6404
11.8286 | -91.6405
11.8289 | -91.6405
11.8292 | -91.6404
11.8295 | -91.641
11.8299 | -91.6407
11.8302 | -91.6399
11.8305 | -91.6395
11.8309 | -91.6397
11.8312 | -91.64
11.8315 | -91.6402
11.8318 | -91.6404
11.8322 | -91.6407
11.8325 | -91.6403
11.8328 | -91.6405
11.8332 | -91.6406
11.8335 | -91.6406
11.8338 | -91.6405
11.8342 | -91.6404
11.8345 | -91.6409
11.8348 | -91.6414
11.8351 | -91.6412
11.8355 | -91.6412
11.8358 | -91.6412
11.8361 | -91.6416
11.8365 | -91.6416
11.8368 | -91.6414
11.8371 | -91.6411
11.8374 | -91.6415
11.8378 | -91.6418
11.8381 | -91.6419
11.8384 | -91.6419
11.8388 | -91.6421
11.8391 | -91.6417
11.8394 | -91.6408
11.8397 | -91.6402
11.8401 | -91.6401
11.8404 | -91.6401
11.8407 | -91.64
11.8411 | -91.6397
11.8414 | -91.6394
11.8417 | -91.6394
11.8421 | -91.639
11.8424 | -91.6387
11.8427 | -91.6385
11.843 | -91.6386
11.8434 | -91.6387
11.8437 | -91.6386
11.844 | -91.6382
11.8444 | -91.6376
11.8447 | -91.6376
11.845 | -91.6375
11.8453 | -91.6376
11.8457 | -91.6379
11.846 | -91.6379
11.8463 | -91.6378
11.8467 | -91.6378
11.847 | -91.6378
11.8473 | -91.637
11.8476 | -91.6372
11.848 | -91.6375
11.8483 | -91.6375
11.8486 | -91.6372
11.849 | -91.6369
11.8493 | -91.636
11.8496 | -91.6358
11.85 | -91.6356
11.8503 | -91.6356
11.8506 | -91.6357
11.8509 | -91.6356
11.8513 | -91.6352
11.8516 | -91.6344
11.8519 | -91.6342
11.8523 | -91.6341
11.8526 | -91.6338
11.8529 | -91.6341
11.8532 | -91.6342
11.8536 | -91.635
11.8539 | -91.636
11.8542 | -91.6353
11.8546 | -91.6357
11.8549 | -91.6362
11.8552 | -91.636
11.8555 | -91.6351
11.8559 | -91.6345
11.8562 | -91.6342
11.8565 | -91.634
11.8569 | -91.6339
11.8572 | -91.634
11.8575 | -91.6337
11.8579 | -91.6339
11.8582 | -91.6341
11.8585 | -91.6342
11.8588 | -91.6343
11.8592 | -91.6343
11.8595 | -91.6343
11.8598 | -91.6344
11.8602 | -91.6349
11.8605 | -91.6346
11.8608 | -91.6343
11.8611 | -91.6341
11.8615 | -91.6346
11.8618 | -91.6345
11.8621 | -91.6339
11.8625 | -91.6334
11.8628 | -91.632
11.8631 | -91.6316
11.8634 | -91.6306
11.8638 | -91.6287
11.8641 | -91.6286
11.8644 | -91.6285
11.8648 | -91.6285
11.8651 | -91.6283
11.8654 | -91.6279
11.8658 | -91.6276
11.8661 | -91.6275
11.8664 | -91.6275
11.8667 | -91.6269
11.8671 | -91.6267
11.8674 | -91.6266
11.8677 | -91.6267
11.8681 | -91.6265
11.8684 | -91.626
11.8687 | -91.6262
11.869 | -91.6265
11.8694 | -91.6266
11.8697 | -91.6271
11.87 | -91.6273
11.8704 | -91.6265
11.8707 | -91.6265
11.871 | -91.6262
11.8713 | -91.6261
11.8717 | -91.6261
11.872 | -91.6262
11.8723 | -91.626
11.8727 | -91.6257
11.873 | -91.6254
11.8733 | -91.6253
11.8737 | -91.6251
11.874 | -91.6248
11.8743 | -91.6244
11.8746 | -91.6242
11.875 | -91.624
11.8753 | -91.6236
11.8756 | -91.6234
11.876 | -91.6231
11.8763 | -91.623
11.8766 | -91.623
11.8769 | -91.6226
11.8773 | -91.6224
11.8776 | -91.6223
11.8779 | -91.6222
11.8783 | -91.6222
11.8786 | -91.622
11.8789 | -91.6215
11.8792 | -91.6215
11.8796 | -91.6213
11.8799 | -91.6206
11.8802 | -91.6205
11.8806 | -91.6203
11.8809 | -91.6202
11.8812 | -91.6197
11.8816 | -91.6199
11.8819 | -91.6204
11.8822 | -91.6205
11.8825 | -91.6204
11.8829 | -91.6199
11.8832 | -91.6196
11.8835 | -91.6186
11.8839 | -91.6184
11.8842 | -91.6182
11.8845 | -91.6177
11.8848 | -91.6174
11.8852 | -91.6174
11.8855 | -91.6174
11.8858 | -91.6173
11.8862 | -91.6174
11.8865 | -91.6171
11.8868 | -91.6158
11.8871 | -91.6154
11.8875 | -91.6156
11.8878 | -91.6157
11.8881 | -91.6156
11.8885 | -91.6159
11.8888 | -91.6163
11.8891 | -91.6162
11.8894 | -91.6161
11.8898 | -91.6153
11.8901 | -91.6146
11.8904 | -91.6145
11.8908 | -91.6143
11.8911 | -91.6143
11.8914 | -91.6141
11.8918 | -91.6141
11.8921 | -91.6135
11.8924 | -91.6134
11.8927 | -91.613
11.8931 | -91.6129
11.8934 | -91.6128
11.8937 | -91.6126
11.8941 | -91.6119
11.8944 | -91.6112
11.8947 | -91.611
11.895 | -91.6106
11.8954 | -91.6103
11.8957 | -91.6102
11.896 | -91.6099
11.8964 | -91.6099
11.8967 | -91.61
11.897 | -91.6109
11.8973 | -91.6109
11.8977 | -91.6109
11.898 | -91.6108
11.8983 | -91.6105
11.8987 | -91.6101
11.899 | -91.6098
11.8993 | -91.6095
11.8997 | -91.6094
11.9 | -91.6093
11.9003 | -91.6086
11.9006 | -91.6083
11.901 | -91.6081
11.9013 | -91.6078
11.9016 | -91.6075
11.902 | -91.6077
11.9023 | -91.6073
11.9026 | -91.6071
11.9029 | -91.6066
11.9033 | -91.6049
11.9036 | -91.6036
11.9039 | -91.6032
11.9043 | -91.6033
11.9046 | -91.6032
11.9049 | -91.6027
11.9052 | -91.603
11.9056 | -91.603
11.9059 | -91.6025
11.9062 | -91.6024
11.9066 | -91.6022
11.9069 | -91.6023
11.9072 | -91.6028
11.9076 | -91.6027
11.9079 | -91.6026
11.9082 | -91.6025
11.9085 | -91.6023
11.9089 | -91.6021
11.9092 | -91.602
11.9095 | -91.6016
11.9099 | -91.6017
11.9102 | -91.6017
11.9105 | -91.6015
11.9108 | -91.6014
11.9112 | -91.6011
11.9115 | -91.6003
11.9118 | -91.5998
11.9122 | -91.5993
11.9125 | -91.599
11.9128 | -91.5987
11.9131 | -91.5986
11.9135 | -91.5985
11.9138 | -91.5984
11.9141 | -91.5984
11.9145 | -91.5983
11.9148 | -91.5979
11.9151 | -91.5971
11.9155 | -91.5966
11.9158 | -91.5965
11.9161 | -91.5968
11.9164 | -91.5975
11.9168 | -91.5979
11.9171 | -91.5975
11.9174 | -91.5976
11.9178 | -91.597
11.9181 | -91.5973
11.9184 | -91.5974
11.9187 | -91.5975
11.9191 | -91.5973
11.9194 | -91.5975
11.9197 | -91.5972
11.9201 | -91.597
11.9204 | -91.597
11.9207 | -91.5973
11.921 | -91.5976
11.9214 | -91.5977
11.9217 | -91.5979
11.922 | -91.5974
11.9224 | -91.596
11.9227 | -91.5949
11.923 | -91.5949
11.9233 | -91.5952
11.9237 | -91.5952
11.924 | -91.5951
11.9243 | -91.5949
11.9247 | -91.5948
11.925 | -91.5946
11.9253 | -91.5945
11.9257 | -91.5944
11.926 | -91.5941
11.9263 | -91.5935
11.9266 | -91.5934
11.927 | -91.5932
11.9273 | -91.5931
11.9276 | -91.5932
11.928 | -91.5931
11.9283 | -91.5928
11.9286 | -91.593
11.9289 | -91.593
11.9293 | -91.5926
11.9296 | -91.5928
11.9299 | -91.5926
11.9303 | -91.5926
11.9306 | -91.5928
11.9309 | -91.5927
11.9313 | -91.5927
11.9316 | -91.5929
11.9319 | -91.5932
11.9322 | -91.5934
11.9326 | -91.5934
11.9329 | -91.5932
11.9332 | -91.5929
11.9336 | -91.5921
11.9339 | -91.5913
11.9342 | -91.5908
11.9345 | -91.5903
11.9349 | -91.5899
11.9352 | -91.5897
11.9355 | -91.5898
11.9359 | -91.5901
11.9362 | -91.5903
11.9365 | -91.59
11.9368 | -91.5901
11.9372 | -91.5901
11.9375 | -91.5899
11.9378 | -91.5897
11.9382 | -91.5897
11.9385 | -91.5898
11.9388 | -91.5896
11.9391 | -91.5895
11.9395 | -91.5897
11.9398 | -91.5896
11.9401 | -91.5894
11.9405 | -91.5892
11.9408 | -91.5886
11.9411 | -91.588
11.9415 | -91.5877
11.9418 | -91.5875
11.9421 | -91.5867
11.9424 | -91.5864
11.9428 | -91.5863
11.9431 | -91.586
11.9434 | -91.5856
11.9438 | -91.5852
11.9441 | -91.5849
11.9444 | -91.5847
11.9447 | -91.5846
11.9451 | -91.5843
11.9454 | -91.5839
11.9457 | -91.5832
11.9461 | -91.5829
11.9464 | -91.5826
11.9467 | -91.5824
11.947 | -91.582
11.9474 | -91.5812
11.9477 | -91.581
11.948 | -91.5809
11.9484 | -91.5807
11.9487 | -91.5803
11.949 | -91.5798
11.9494 | -91.5795
11.9497 | -91.5792
11.95 | -91.579
11.9503 | -91.5791
11.9507 | -91.5786
11.951 | -91.578
11.9513 | -91.5779
11.9517 | -91.5779
11.952 | -91.5776
11.9523 | -91.5774
11.9526 | -91.5769
11.953 | -91.5769
11.9533 | -91.5767
11.9536 | -91.5765
11.954 | -91.5763
11.9543 | -91.5758
11.9546 | -91.5754
11.9549 | -91.5758
11.9553 | -91.5756
11.9556 | -91.5755
11.9559 | -91.5752
11.9563 | -91.5748
11.9566 | -91.5749
11.9569 | -91.5748
11.9573 | -91.5745
11.9576 | -91.5742
11.9579 | -91.574
11.9582 | -91.5731
11.9586 | -91.5721
11.9589 | -91.572
11.9592 | -91.5718
11.9596 | -91.5719
11.9599 | -91.5716
11.9602 | -91.5714
11.9605 | -91.5712
11.9609 | -91.571
11.9612 | -91.571
11.9615 | -91.5715
11.9619 | -91.5716
11.9622 | -91.5719
11.9625 | -91.5724
11.9628 | -91.5721
11.9632 | -91.5718
11.9635 | -91.5713
11.9638 | -91.5706
11.9642 | -91.5702
11.9645 | -91.57
11.9648 | -91.57
11.9652 | -91.57
11.9655 | -91.5701
11.9658 | -91.57
11.9661 | -91.5696
11.9665 | -91.5692
11.9668 | -91.5693
11.9671 | -91.5689
11.9675 | -91.5687
11.9678 | -91.5685
11.9681 | -91.5684
11.9684 | -91.5682
11.9688 | -91.5682
11.9691 | -91.5682
11.9694 | -91.5682
11.9698 | -91.568
11.9701 | -91.5679
11.9704 | -91.5677
11.9707 | -91.5674
11.9711 | -91.5672
11.9714 | -91.5672
11.9717 | -91.5669
11.9721 | -91.5665
11.9724 | -91.5663
11.9727 | -91.566
11.9731 | -91.5658
11.9734 | -91.5656
11.9737 | -91.5653
11.974 | -91.5652
11.9744 | -91.5649
11.9747 | -91.5649
11.975 | -91.5647
11.9754 | -91.5643
11.9757 | -91.5638
11.976 | -91.5635
11.9763 | -91.5632
11.9767 | -91.5629
11.977 | -91.5624
11.9773 | -91.5617
11.9777 | -91.5611
11.978 | -91.561
11.9783 | -91.5607
11.9786 | -91.5603
11.979 | -91.5601
11.9793 | -91.56
11.9796 | -91.5597
11.98 | -91.5595
11.9803 | -91.5593
11.9806 | -91.5593
11.9809 | -91.5595
11.9813 | -91.5595
11.9816 | -91.5592
11.9819 | -91.559
11.9823 | -91.5589
11.9826 | -91.5585
11.9829 | -91.5582
11.9833 | -91.5579
11.9836 | -91.5579
11.9839 | -91.5575
11.9842 | -91.5574
11.9846 | -91.557
11.9849 | -91.5568
11.9852 | -91.5573
11.9856 | -91.5577
11.9859 | -91.5578
11.9862 | -91.5572
11.9865 | -91.557
11.9869 | -91.5569
11.9872 | -91.557
11.9875 | -91.557
11.9879 | -91.5569
11.9882 | -91.5567
11.9885 | -91.5562
11.9888 | -91.5553
11.9892 | -91.5546
11.9895 | -91.5542
11.9898 | -91.5538
11.9902 | -91.5532
11.9905 | -91.553
11.9908 | -91.5531
11.9912 | -91.5526
11.9915 | -91.5527
11.9918 | -91.5527
11.9921 | -91.5523
11.9925 | -91.552
11.9928 | -91.5517
11.9931 | -91.5514
11.9935 | -91.5512
11.9938 | -91.5511
11.9941 | -91.5513
11.9944 | -91.5516
11.9948 | -91.5515
11.9951 | -91.5513
11.9954 | -91.5506
11.9958 | -91.5502
11.9961 | -91.5502
11.9964 | -91.5501
11.9967 | -91.5494
11.9971 | -91.5491
11.9974 | -91.5487
11.9977 | -91.5483
11.9981 | -91.5481
11.9984 | -91.5477
11.9987 | -91.5474
11.9991 | -91.5467
11.9994 | -91.5463
11.9997 | -91.5464
12 | -91.5464
12.0004 | -91.5462
12.0007 | -91.5458
12.001 | -91.5454
12.0014 | -91.5453
12.0017 | -91.5454
12.002 | -91.5453
12.0023 | -91.5451
12.0027 | -91.5444
12.003 | -91.5445
12.0033 | -91.5444
12.0037 | -91.5445
12.004 | -91.5445
12.0043 | -91.5441
12.0046 | -91.5438
12.005 | -91.5437
12.0053 | -91.5436
12.0056 | -91.5432
12.006 | -91.5428
12.0063 | -91.5425
12.0066 | -91.5423
12.007 | -91.5419
12.0073 | -91.5416
12.0076 | -91.5413
12.0079 | -91.5407
12.0083 | -91.5404
12.0086 | -91.5401
12.0089 | -91.5395
12.0093 | -91.5392
12.0096 | -91.5387
12.0099 | -91.5385
12.0102 | -91.5386
12.0106 | -91.5384
12.0109 | -91.5381
12.0112 | -91.5379
12.0116 | -91.5381
12.0119 | -91.5381
12.0122 | -91.5378
12.0125 | -91.5377
12.0129 | -91.5374
12.0132 | -91.5369
12.0135 | -91.5365
12.0139 | -91.5364
12.0142 | -91.5361
12.0145 | -91.5357
12.0149 | -91.5347
12.0152 | -91.534
12.0155 | -91.5337
12.0158 | -91.5335
12.0162 | -91.5328
12.0165 | -91.533
12.0168 | -91.5327
12.0172 | -91.532
12.0175 | -91.5318
12.0178 | -91.5313
12.0181 | -91.531
12.0185 | -91.5307
12.0188 | -91.5306
12.0191 | -91.5303
12.0195 | -91.5299
12.0198 | -91.5293
12.0201 | -91.5291
12.0204 | -91.5286
12.0208 | -91.5284
12.0211 | -91.5282
12.0214 | -91.5281
12.0218 | -91.5277
12.0221 | -91.5271
12.0224 | -91.5267
12.0228 | -91.5266
12.0231 | -91.5266
12.0234 | -91.5263
12.0237 | -91.5258
12.0241 | -91.5257
12.0244 | -91.5258
12.0247 | -91.5256
12.0251 | -91.5252
12.0254 | -91.5252
12.0257 | -91.525
12.026 | -91.5241
12.0264 | -91.5237
12.0267 | -91.5234
12.027 | -91.5231
12.0274 | -91.5227
12.0277 | -91.5221
12.028 | -91.5215
12.0283 | -91.5212
12.0287 | -91.5206
12.029 | -91.52
12.0293 | -91.5195
12.0297 | -91.5198
12.03 | -91.5197
12.0303 | -91.5193
12.0306 | -91.5191
12.031 | -91.5189
12.0313 | -91.5185
12.0316 | -91.5182
12.032 | -91.5177
12.0323 | -91.5169
12.0326 | -91.5168
12.033 | -91.5165
12.0333 | -91.5162
12.0336 | -91.5161
12.0339 | -91.5159
12.0343 | -91.5157
12.0346 | -91.5154
12.0349 | -91.5153
12.0353 | -91.5149
12.0356 | -91.514
12.0359 | -91.5135
12.0362 | -91.513
12.0366 | -91.5127
12.0369 | -91.5124
12.0372 | -91.5121
12.0376 | -91.5119
12.0379 | -91.5118
12.0382 | -91.5116
12.0385 | -91.5114
12.0389 | -91.5112
12.0392 | -91.511
12.0395 | -91.5106
12.0399 | -91.5103
12.0402 | -91.5101
12.0405 | -91.5098
12.0409 | -91.5094
12.0412 | -91.5085
12.0415 | -91.5083
12.0418 | -91.5085
12.0422 | -91.5083
12.0425 | -91.5081
12.0428 | -91.5079
12.0432 | -91.5078
12.0435 | -91.5074
12.0438 | -91.5071
12.0441 | -91.5069
12.0445 | -91.5066
12.0448 | -91.5064
12.0451 | -91.5061
12.0455 | -91.5063
12.0458 | -91.5059
12.0461 | -91.5055
12.0464 | -91.5052
12.0468 | -91.505
12.0471 | -91.5049
12.0474 | -91.5047
12.0478 | -91.5045
12.0481 | -91.5043
12.0484 | -91.5039
12.0488 | -91.5033
12.0491 | -91.5027
12.0494 | -91.5024
12.0497 | -91.502
12.0501 | -91.5018
12.0504 | -91.5016
12.0507 | -91.5013
12.0511 | -91.501
12.0514 | -91.5004
12.0517 | -91.4999
12.052 | -91.4995
12.0524 | -91.4992
12.0527 | -91.4992
12.053 | -91.4992
12.0534 | -91.4989
12.0537 | -91.4985
12.054 | -91.4981
12.0543 | -91.4976
12.0547 | -91.4978
12.055 | -91.4976
12.0553 | -91.497
12.0557 | -91.4968
12.056 | -91.4964
12.0563 | -91.4962
12.0567 | -91.496
12.057 | -91.4957
12.0573 | -91.4958
12.0576 | -91.4956
12.058 | -91.4953
12.0583 | -91.495
12.0586 | -91.4942
12.059 | -91.4937
12.0593 | -91.4929
12.0596 | -91.4921
12.0599 | -91.4914
12.0603 | -91.4905
12.0606 | -91.4899
12.0609 | -91.4899
12.0613 | -91.4897
12.0616 | -91.4893
12.0619 | -91.4889
12.0622 | -91.4884
12.0626 | -91.4884
12.0629 | -91.4881
12.0632 | -91.4876
12.0636 | -91.4872
12.0639 | -91.4868
12.0642 | -91.4861
12.0646 | -91.4857
12.0649 | -91.4856
12.0652 | -91.4851
12.0655 | -91.4846
12.0659 | -91.4842
12.0662 | -91.4843
12.0665 | -91.4842
12.0669 | -91.4843
12.0672 | -91.4845
12.0675 | -91.4845
12.0678 | -91.4846
12.0682 | -91.4844
12.0685 | -91.4841
12.0688 | -91.4836
12.0692 | -91.4831
12.0695 | -91.4829
12.0698 | -91.4825
12.0701 | -91.482
12.0705 | -91.4814
12.0708 | -91.4808
12.0711 | -91.4804
12.0715 | -91.48
12.0718 | -91.4795
12.0721 | -91.4791
12.0724 | -91.4788
12.0728 | -91.4782
12.0731 | -91.4775
12.0734 | -91.4766
12.0738 | -91.4757
12.0741 | -91.4753
12.0744 | -91.4746
12.0748 | -91.4739
12.0751 | -91.4738
12.0754 | -91.4715
12.0757 | -91.4707
12.0761 | -91.4707
12.0764 | -91.47
12.0767 | -91.4688
12.0771 | -91.4664
12.0774 | -91.4654
12.0777 | -91.465
12.078 | -91.464
12.0784 | -91.4639
12.0787 | -91.464
12.079 | -91.4631
12.0794 | -91.4606
12.0797 | -91.459
12.08 | -91.4569
12.0803 | -91.4563
12.0807 | -91.456
12.081 | -91.4556
12.0813 | -91.4549
12.0817 | -91.4538
12.082 | -91.453
12.0823 | -91.4524
12.0827 | -91.4519
12.083 | -91.4511
12.0833 | -91.4508
12.0836 | -91.4508
12.084 | -91.4502
12.0843 | -91.4499
12.0846 | -91.4494
12.085 | -91.4492
12.0853 | -91.4488
12.0856 | -91.4484
12.0859 | -91.4481
12.0863 | -91.4484
12.0866 | -91.4484
12.0869 | -91.4479
12.0873 | -91.4474
12.0876 | -91.4468
12.0879 | -91.4463
12.0882 | -91.4457
12.0886 | -91.445
12.0889 | -91.4441
12.0892 | -91.4432
12.0896 | -91.4429
12.0899 | -91.4429
12.0902 | -91.4424
12.0906 | -91.4418
12.0909 | -91.4418
12.0912 | -91.4418
12.0915 | -91.4414
12.0919 | -91.4411
12.0922 | -91.4411
12.0925 | -91.4409
12.0929 | -91.4405
12.0932 | -91.4398
12.0935 | -91.4391
12.0938 | -91.4383
12.0942 | -91.4375
12.0945 | -91.4371
12.0948 | -91.4366
12.0952 | -91.4362
12.0955 | -91.4357
12.0958 | -91.4354
12.0961 | -91.4349
12.0965 | -91.4343
12.0968 | -91.4341
12.0971 | -91.4337
12.0975 | -91.4333
12.0978 | -91.433
12.0981 | -91.4326
12.0985 | -91.4324
12.0988 | -91.432
12.0991 | -91.4314
12.0994 | -91.4306
12.0998 | -91.43
12.1001 | -91.4293
12.1004 | -91.4285
12.1008 | -91.4277
12.1011 | -91.4273
12.1014 | -91.4269
12.1017 | -91.4267
12.1021 | -91.4263
12.1024 | -91.426
12.1027 | -91.4265
12.1031 | -91.4266
12.1034 | -91.4262
12.1037 | -91.4258
12.104 | -91.4254
12.1044 | -91.4249
12.1047 | -91.4247
12.105 | -91.4242
12.1054 | -91.4236
12.1057 | -91.4232
12.106 | -91.4226
12.1064 | -91.422
12.1067 | -91.4214
12.107 | -91.421
12.1073 | -91.4208
12.1077 | -91.42
12.108 | -91.4197
12.1083 | -91.4192
12.1087 | -91.419
12.109 | -91.4192
12.1093 | -91.4194
12.1096 | -91.4192
12.11 | -91.4191
12.1103 | -91.419
12.1106 | -91.4187
12.111 | -91.4181
12.1113 | -91.4181
12.1116 | -91.4181
12.1119 | -91.4177
12.1123 | -91.4174
12.1126 | -91.4169
12.1129 | -91.4164
12.1133 | -91.4162
12.1136 | -91.416
12.1139 | -91.4156
12.1143 | -91.4151
12.1146 | -91.4151
12.1149 | -91.4148
12.1152 | -91.4144
12.1156 | -91.4137
12.1159 | -91.413
12.1162 | -91.4124
12.1166 | -91.412
12.1169 | -91.4115
12.1172 | -91.4108
12.1175 | -91.4104
12.1179 | -91.4101
12.1182 | -91.4096
12.1185 | -91.4091
12.1189 | -91.4085
12.1192 | -91.4079
12.1195 | -91.4074
12.1198 | -91.4073
12.1202 | -91.407
12.1205 | -91.4066
12.1208 | -91.4054
12.1212 | -91.4051
12.1215 | -91.4054
12.1218 | -91.4052
12.1221 | -91.4049
12.1225 | -91.4045
12.1228 | -91.4042
12.1231 | -91.404
12.1235 | -91.4036
12.1238 | -91.4036
12.1241 | -91.4037
12.1245 | -91.4029
12.1248 | -91.4025
12.1251 | -91.4023
12.1254 | -91.4018
12.1258 | -91.4013
12.1261 | -91.401
12.1264 | -91.4004
12.1268 | -91.4002
12.1271 | -91.3999
12.1274 | -91.3993
12.1277 | -91.3986
12.1281 | -91.3978
12.1284 | -91.3972
12.1287 | -91.3969
12.1291 | -91.3964
12.1294 | -91.3954
12.1297 | -91.3944
12.13 | -91.3939
12.1304 | -91.3937
12.1307 | -91.3937
12.131 | -91.3934
12.1314 | -91.3929
12.1317 | -91.3925
12.132 | -91.3922
12.1324 | -91.3917
12.1327 | -91.3913
12.133 | -91.3904
12.1333 | -91.3897
12.1337 | -91.3893
12.134 | -91.3887
12.1343 | -91.388
12.1347 | -91.3875
12.135 | -91.3872
12.1353 | -91.387
12.1356 | -91.3866
12.136 | -91.3862
12.1363 | -91.3848
12.1366 | -91.3842
12.137 | -91.3841
12.1373 | -91.3838
12.1376 | -91.3841
12.1379 | -91.3841
12.1383 | -91.3839
12.1386 | -91.3835
12.1389 | -91.3832
12.1393 | -91.3828
12.1396 | -91.3823
12.1399 | -91.3817
12.1403 | -91.3808
12.1406 | -91.3805
12.1409 | -91.3804
12.1412 | -91.3801
12.1416 | -91.3798
12.1419 | -91.3794
12.1422 | -91.379
12.1426 | -91.3785
12.1429 | -91.3779
12.1432 | -91.3776
12.1435 | -91.3777
12.1439 | -91.3774
12.1442 | -91.3767
12.1445 | -91.3765
12.1449 | -91.3759
12.1452 | -91.3753
12.1455 | -91.3748
12.1458 | -91.374
12.1462 | -91.3737
12.1465 | -91.3736
12.1468 | -91.3733
12.1472 | -91.3725
12.1475 | -91.3717
12.1478 | -91.3701
12.1482 | -91.369
12.1485 | -91.3685
12.1488 | -91.3681
12.1491 | -91.3678
12.1495 | -91.3675
12.1498 | -91.3672
12.1501 | -91.3668
12.1505 | -91.3662
12.1508 | -91.3658
12.1511 | -91.3654
12.1514 | -91.3644
12.1518 | -91.3638
12.1521 | -91.3633
12.1524 | -91.363
12.1528 | -91.3624
12.1531 | -91.3615
12.1534 | -91.3614
12.1537 | -91.3613
12.1541 | -91.361
12.1544 | -91.3608
12.1547 | -91.3606
12.1551 | -91.3604
12.1554 | -91.3602
12.1557 | -91.3597
12.1561 | -91.359
12.1564 | -91.3582
12.1567 | -91.358
12.157 | -91.3579
12.1574 | -91.3577
12.1577 | -91.357
12.158 | -91.3567
12.1584 | -91.3563
12.1587 | -91.356
12.159 | -91.3554
12.1593 | -91.3547
12.1597 | -91.3544
12.16 | -91.3542
12.1603 | -91.3544
12.1607 | -91.3544
12.161 | -91.354
12.1613 | -91.3533
12.1616 | -91.3529
12.162 | -91.3532
12.1623 | -91.3534
12.1626 | -91.3531
12.163 | -91.353
12.1633 | -91.3527
12.1636 | -91.3521
12.1639 | -91.3516
12.1643 | -91.3513
12.1646 | -91.3512
12.1649 | -91.3507
12.1653 | -91.3502
12.1656 | -91.3496
12.1659 | -91.3491
12.1663 | -91.3487
12.1666 | -91.3482
12.1669 | -91.348
12.1672 | -91.3477
12.1676 | -91.3475
12.1679 | -91.3473
12.1682 | -91.3469
12.1686 | -91.3463
12.1689 | -91.3457
12.1692 | -91.3451
12.1695 | -91.3446
12.1699 | -91.3439
12.1702 | -91.3434
12.1705 | -91.3429
12.1709 | -91.3423
12.1712 | -91.3415
12.1715 | -91.341
12.1718 | -91.3408
12.1722 | -91.3403
12.1725 | -91.3399
12.1728 | -91.3394
12.1732 | -91.3388
12.1735 | -91.3383
12.1738 | -91.3378
12.1742 | -91.3375
12.1745 | -91.3374
12.1748 | -91.3376
12.1751 | -91.3374
12.1755 | -91.337
12.1758 | -91.3367
12.1761 | -91.3365
12.1765 | -91.3362
12.1768 | -91.3358
12.1771 | -91.3356
12.1774 | -91.3352
12.1778 | -91.3346
12.1781 | -91.334
12.1784 | -91.3335
12.1788 | -91.3328
12.1791 | -91.3323
12.1794 | -91.3319
12.1797 | -91.3315
12.1801 | -91.3312
12.1804 | -91.3309
12.1807 | -91.3305
12.1811 | -91.3303
12.1814 | -91.3299
12.1817 | -91.3295
12.1821 | -91.3291
12.1824 | -91.3287
12.1827 | -91.3284
12.183 | -91.3283
12.1834 | -91.3284
12.1837 | -91.3283
12.184 | -91.3281
12.1844 | -91.3273
12.1847 | -91.3268
12.185 | -91.3264
12.1853 | -91.326
12.1857 | -91.3254
12.186 | -91.3248
12.1863 | -91.3242
12.1867 | -91.3235
12.187 | -91.3228
12.1873 | -91.3225
12.1876 | -91.3223
12.188 | -91.3219
12.1883 | -91.3215
12.1886 | -91.3211
12.189 | -91.3206
12.1893 | -91.3202
12.1896 | -91.3197
12.19 | -91.3193
12.1903 | -91.3186
12.1906 | -91.3181
12.1909 | -91.3178
12.1913 | -91.3173
12.1916 | -91.3168
12.1919 | -91.3162
12.1923 | -91.3156
12.1926 | -91.3149
12.1929 | -91.3143
12.1932 | -91.3143
12.1936 | -91.3139
12.1939 | -91.3135
12.1942 | -91.313
12.1946 | -91.3123
12.1949 | -91.3118
12.1952 | -91.3115
12.1955 | -91.3112
12.1959 | -91.3106
12.1962 | -91.3101
12.1965 | -91.3098
12.1969 | -91.3094
12.1972 | -91.309
12.1975 | -91.3086
12.1979 | -91.3083
12.1982 | -91.3084
12.1985 | -91.3082
12.1988 | -91.3081
12.1992 | -91.3078
12.1995 | -91.3075
12.1998 | -91.3072
12.2002 | -91.3068
12.2005 | -91.3063
12.2008 | -91.3056
12.2011 | -91.3049
12.2015 | -91.3044
12.2018 | -91.3042
12.2021 | -91.304
12.2025 | -91.3033
12.2028 | -91.3026
12.2031 | -91.3021
12.2034 | -91.3017
12.2038 | -91.301
12.2041 | -91.3006
12.2044 | -91.3003
12.2048 | -91.3003
12.2051 | -91.2998
12.2054 | -91.2992
12.2058 | -91.2989
12.2061 | -91.2987
12.2064 | -91.2982
12.2067 | -91.2978
12.2071 | -91.2971
12.2074 | -91.2965
12.2077 | -91.2963
12.2081 | -91.2958
12.2084 | -91.2954
12.2087 | -91.2946
12.209 | -91.2936
12.2094 | -91.293
12.2097 | -91.2923
12.21 | -91.2915
12.2104 | -91.2908
12.2107 | -91.2901
12.211 | -91.2898
12.2113 | -91.2895
12.2117 | -91.289
12.212 | -91.289
12.2123 | -91.2889
12.2127 | -91.2884
12.213 | -91.2876
12.2133 | -91.2871
12.2136 | -91.2869
12.214 | -91.2865
12.2143 | -91.286
12.2146 | -91.2857
12.215 | -91.2856
12.2153 | -91.2855
12.2156 | -91.2851
12.216 | -91.2846
12.2163 | -91.2841
12.2166 | -91.2833
12.2169 | -91.2824
12.2173 | -91.2817
12.2176 | -91.2813
12.2179 | -91.2813
12.2183 | -91.2808
12.2186 | -91.2805
12.2189 | -91.2805
12.2192 | -91.2799
12.2196 | -91.2795
12.2199 | -91.2791
12.2202 | -91.2787
12.2206 | -91.2783
12.2209 | -91.2783
12.2212 | -91.2784
12.2215 | -91.2779
12.2219 | -91.2776
12.2222 | -91.2772
12.2225 | -91.2769
12.2229 | -91.2766
12.2232 | -91.2762
12.2235 | -91.2759
12.2239 | -91.2755
12.2242 | -91.2751
12.2245 | -91.2749
12.2248 | -91.2746
12.2252 | -91.2741
12.2255 | -91.2736
12.2258 | -91.2733
12.2262 | -91.2733
12.2265 | -91.2725
12.2268 | -91.2721
12.2271 | -91.2718
12.2275 | -91.2712
12.2278 | -91.2708
12.2281 | -91.2706
12.2285 | -91.2705
12.2288 | -91.2703
12.2291 | -91.2703
12.2294 | -91.2708
12.2298 | -91.2712
12.2301 | -91.2712
12.2304 | -91.271
12.2308 | -91.271
12.2311 | -91.2711
12.2314 | -91.2705
12.2318 | -91.2697
12.2321 | -91.2692
12.2324 | -91.2688
12.2327 | -91.2684
12.2331 | -91.2681
12.2334 | -91.2678
12.2337 | -91.2673
12.2341 | -91.2671
12.2344 | -91.2668
12.2347 | -91.2665
12.235 | -91.2664
12.2354 | -91.2661
12.2357 | -91.2655
12.236 | -91.2649
12.2364 | -91.2644
12.2367 | -91.2643
12.237 | -91.2639
12.2373 | -91.2634
12.2377 | -91.263
12.238 | -91.263
12.2383 | -91.2627
12.2387 | -91.2624
12.239 | -91.2619
12.2393 | -91.2616
12.2397 | -91.2612
12.24 | -91.2608
12.2403 | -91.2605
12.2406 | -91.2606
12.241 | -91.2602
12.2413 | -91.259
12.2416 | -91.2583
12.242 | -91.2576
12.2423 | -91.2567
12.2426 | -91.2565
12.2429 | -91.2564
12.2433 | -91.2561
12.2436 | -91.2559
12.2439 | -91.2557
12.2443 | -91.2556
12.2446 | -91.2553
12.2449 | -91.255
12.2452 | -91.2548
12.2456 | -91.2549
12.2459 | -91.2547
12.2462 | -91.2541
12.2466 | -91.2535
12.2469 | -91.253
12.2472 | -91.2525
12.2476 | -91.2522
12.2479 | -91.2518
12.2482 | -91.2514
12.2485 | -91.2511
12.2489 | -91.2508
12.2492 | -91.2504
12.2495 | -91.2499
12.2499 | -91.2497
12.2502 | -91.2496
12.2505 | -91.2495
12.2508 | -91.2493
12.2512 | -91.2493
12.2515 | -91.2492
12.2518 | -91.249
12.2522 | -91.2491
12.2525 | -91.2492
12.2528 | -91.2488
12.2531 | -91.2483
12.2535 | -91.2483
12.2538 | -91.2481
12.2541 | -91.2478
12.2545 | -91.2475
12.2548 | -91.2474
12.2551 | -91.2473
12.2555 | -91.247
12.2558 | -91.2468
12.2561 | -91.2467
12.2564 | -91.2464
12.2568 | -91.2458
12.2571 | -91.2452
12.2574 | -91.2451
12.2578 | -91.2449
12.2581 | -91.2447
12.2584 | -91.2444
12.2587 | -91.2443
12.2591 | -91.2443
12.2594 | -91.2441
12.2597 | -91.244
12.2601 | -91.2434
12.2604 | -91.2424
12.2607 | -91.242
12.261 | -91.2416
12.2614 | -91.2413
12.2617 | -91.2411
12.262 | -91.2408
12.2624 | -91.2405
12.2627 | -91.2404
12.263 | -91.241
12.2634 | -91.2413
12.2637 | -91.2411
12.264 | -91.2409
12.2643 | -91.2412
12.2647 | -91.2412
12.265 | -91.2411
12.2653 | -91.241
12.2657 | -91.2408
12.266 | -91.2406
12.2663 | -91.2406
12.2666 | -91.2402
12.267 | -91.24
12.2673 | -91.2405
12.2676 | -91.2406
12.268 | -91.2407
12.2683 | -91.2399
12.2686 | -91.2389
12.2689 | -91.2387
12.2693 | -91.2383
12.2696 | -91.238
12.2699 | -91.2376
12.2703 | -91.2371
12.2706 | -91.2368
12.2709 | -91.2368
12.2712 | -91.2367
12.2716 | -91.2364
12.2719 | -91.236
12.2722 | -91.2357
12.2726 | -91.2352
12.2729 | -91.235
12.2732 | -91.2347
12.2736 | -91.2343
12.2739 | -91.2336
12.2742 | -91.2336
12.2745 | -91.2337
12.2749 | -91.2335
12.2752 | -91.2334
12.2755 | -91.2333
12.2759 | -91.2331
12.2762 | -91.2325
12.2765 | -91.2308
12.2768 | -91.229
12.2772 | -91.2283
12.2775 | -91.228
12.2778 | -91.2278
12.2782 | -91.2273
12.2785 | -91.227
12.2788 | -91.2272
12.2791 | -91.2269
12.2795 | -91.2269
12.2798 | -91.2272
12.2801 | -91.2271
12.2805 | -91.2269
12.2808 | -91.2274
12.2811 | -91.2276
12.2815 | -91.2271
12.2818 | -91.2263
12.2821 | -91.2259
12.2824 | -91.2257
12.2828 | -91.225
12.2831 | -91.2249
12.2834 | -91.2252
12.2838 | -91.2255
12.2841 | -91.226
12.2844 | -91.2269
12.2847 | -91.2276
12.2851 | -91.2275
12.2854 | -91.2273
12.2857 | -91.2271
12.2861 | -91.2268
12.2864 | -91.2265
12.2867 | -91.2264
12.287 | -91.2262
12.2874 | -91.2259
12.2877 | -91.2254
12.288 | -91.2254
12.2884 | -91.2257
12.2887 | -91.2256
12.289 | -91.2254
12.2894 | -91.2252
12.2897 | -91.2243
12.29 | -91.2237
12.2903 | -91.2235
12.2907 | -91.2233
12.291 | -91.2226
12.2913 | -91.2222
12.2917 | -91.2216
12.292 | -91.2208
12.2923 | -91.2203
12.2926 | -91.2198
12.293 | -91.2198
12.2933 | -91.22
12.2936 | -91.2198
12.294 | -91.2192
12.2943 | -91.2186
12.2946 | -91.2183
12.2949 | -91.2179
12.2953 | -91.218
12.2956 | -91.2178
12.2959 | -91.2174
12.2963 | -91.2173
12.2966 | -91.2176
12.2969 | -91.2172
12.2973 | -91.217
12.2976 | -91.217
12.2979 | -91.2172
12.2982 | -91.2172
12.2986 | -91.217
12.2989 | -91.2167
12.2992 | -91.2163
12.2996 | -91.2161
12.2999 | -91.2158
12.3002 | -91.2157
12.3005 | -91.2159
12.3009 | -91.2156
12.3012 | -91.2158
12.3015 | -91.216
12.3019 | -91.2155
12.3022 | -91.215
12.3025 | -91.2149
12.3028 | -91.2148
12.3032 | -91.2145
12.3035 | -91.2142
12.3038 | -91.2138
12.3042 | -91.2138
12.3045 | -91.2141
12.3048 | -91.2141
12.3051 | -91.2145
12.3055 | -91.2144
12.3058 | -91.2141
12.3061 | -91.2141
12.3065 | -91.2138
12.3068 | -91.2132
12.3071 | -91.2116
12.3075 | -91.211
12.3078 | -91.2105
12.3081 | -91.2102
12.3084 | -91.2103
12.3088 | -91.2104
12.3091 | -91.2103
12.3094 | -91.2101
12.3098 | -91.2099
12.3101 | -91.2098
12.3104 | -91.2095
12.3107 | -91.209
12.3111 | -91.2088
12.3114 | -91.2091
12.3117 | -91.2089
12.3121 | -91.2088
12.3124 | -91.2085
12.3127 | -91.2087
12.3131 | -91.2086
12.3134 | -91.2085
12.3137 | -91.2085
12.314 | -91.2079
12.3144 | -91.2076
12.3147 | -91.2072
12.315 | -91.2068
12.3154 | -91.2068
12.3157 | -91.2067
12.316 | -91.2067
12.3163 | -91.2065
12.3167 | -91.2058
12.317 | -91.2056
12.3173 | -91.2054
12.3177 | -91.2053
12.318 | -91.2044
12.3183 | -91.2037
12.3186 | -91.2031
12.319 | -91.2023
12.3193 | -91.2017
12.3196 | -91.2015
12.32 | -91.2013
12.3203 | -91.2012
12.3206 | -91.201
12.3209 | -91.2009
12.3213 | -91.2007
12.3216 | -91.2007
12.3219 | -91.2005
12.3223 | -91.2006
12.3226 | -91.2006
12.3229 | -91.2005
12.3233 | -91.2003
12.3236 | -91.2
12.3239 | -91.1995
12.3242 | -91.1993
12.3246 | -91.1995
12.3249 | -91.1995
12.3252 | -91.1994
12.3256 | -91.1992
12.3259 | -91.1992
12.3262 | -91.1992
12.3265 | -91.1987
12.3269 | -91.1985
12.3272 | -91.1991
12.3275 | -91.1996
12.3279 | -91.1996
12.3282 | -91.1994
12.3285 | -91.1993
12.3288 | -91.1993
12.3292 | -91.1989
12.3295 | -91.1983
12.3298 | -91.198
12.3302 | -91.1981
12.3305 | -91.1982
12.3308 | -91.1979
12.3312 | -91.1977
12.3315 | -91.1978
12.3318 | -91.1977
12.3321 | -91.1976
12.3325 | -91.1975
12.3328 | -91.1974
12.3331 | -91.1973
12.3335 | -91.1971
12.3338 | -91.1967
12.3341 | -91.1964
12.3344 | -91.196
12.3348 | -91.1959
12.3351 | -91.1962
12.3354 | -91.1966
12.3358 | -91.1967
12.3361 | -91.1967
12.3364 | -91.1967
12.3367 | -91.196
12.3371 | -91.1953
12.3374 | -91.1953
12.3377 | -91.1952
12.3381 | -91.195
12.3384 | -91.1949
12.3387 | -91.1949
12.3391 | -91.195
12.3394 | -91.1952
12.3397 | -91.1951
12.34 | -91.1948
12.3404 | -91.1946
12.3407 | -91.1942
12.341 | -91.1941
12.3414 | -91.194
12.3417 | -91.1936
12.342 | -91.1931
12.3423 | -91.1923
12.3427 | -91.1918
12.343 | -91.1917
12.3433 | -91.1915
12.3437 | -91.191
12.344 | -91.1907
12.3443 | -91.1906
12.3446 | -91.1906
12.345 | -91.1907
12.3453 | -91.1903
12.3456 | -91.1904
12.346 | -91.1904
12.3463 | -91.1907
12.3466 | -91.1906
12.347 | -91.1902
12.3473 | -91.1898
12.3476 | -91.1897
12.3479 | -91.1895
12.3483 | -91.1895
12.3486 | -91.1895
12.3489 | -91.1893
12.3493 | -91.1892
12.3496 | -91.1891
12.3499 | -91.189
12.3502 | -91.1889
12.3506 | -91.1887
12.3509 | -91.1887
12.3512 | -91.1887
12.3516 | -91.1886
12.3519 | -91.1888
12.3522 | -91.1886
12.3525 | -91.1882
12.3529 | -91.1884
12.3532 | -91.1882
12.3535 | -91.1878
12.3539 | -91.1879
12.3542 | -91.1881
12.3545 | -91.1879
12.3549 | -91.1882
12.3552 | -91.188
12.3555 | -91.188
12.3558 | -91.1881
12.3562 | -91.1881
12.3565 | -91.1883
12.3568 | -91.1882
12.3572 | -91.1882
12.3575 | -91.1882
12.3578 | -91.1879
12.3581 | -91.1876
12.3585 | -91.1876
12.3588 | -91.1877
12.3591 | -91.1879
12.3595 | -91.1882
12.3598 | -91.1883
12.3601 | -91.1883
12.3604 | -91.1886
12.3608 | -91.1887
12.3611 | -91.1888
12.3614 | -91.1889
12.3618 | -91.1889
12.3621 | -91.1885
12.3624 | -91.1884
12.3627 | -91.1884
12.3631 | -91.1883
12.3634 | -91.1881
12.3637 | -91.188
12.3641 | -91.188
12.3644 | -91.1881
12.3647 | -91.1881
12.3651 | -91.188
12.3654 | -91.1878
12.3657 | -91.1879
12.366 | -91.1878
12.3664 | -91.1874
12.3667 | -91.1869
12.367 | -91.1869
12.3674 | -91.187
12.3677 | -91.1871
12.368 | -91.1872
12.3683 | -91.1873
12.3687 | -91.1873
12.369 | -91.1871
12.3693 | -91.1873
12.3697 | -91.1872
12.37 | -91.1871
12.3703 | -91.187
12.3706 | -91.187
12.371 | -91.1869
12.3713 | -91.1868
12.3716 | -91.1868
12.372 | -91.1868
12.3723 | -91.1871
12.3726 | -91.1876
12.373 | -91.1878
12.3733 | -91.188
12.3736 | -91.1878
12.3739 | -91.1873
12.3743 | -91.1872
12.3746 | -91.1879
12.3749 | -91.1881
12.3753 | -91.1881
12.3756 | -91.188
12.3759 | -91.1879
12.3762 | -91.188
12.3766 | -91.1881
12.3769 | -91.188
12.3772 | -91.1881
12.3776 | -91.188
12.3779 | -91.1879
12.3782 | -91.1878
12.3785 | -91.1875
12.3789 | -91.1875
12.3792 | -91.1875
12.3795 | -91.1869
12.3799 | -91.186
12.3802 | -91.186
12.3805 | -91.1858
12.3809 | -91.1857
12.3812 | -91.1857
12.3815 | -91.1858
12.3818 | -91.1859
12.3822 | -91.1856
12.3825 | -91.1857
12.3828 | -91.1856
12.3832 | -91.1859
12.3835 | -91.1855
12.3838 | -91.1847
12.3841 | -91.1845
12.3845 | -91.1846
12.3848 | -91.1846
12.3851 | -91.1838
12.3855 | -91.1823
12.3858 | -91.1817
12.3861 | -91.1813
12.3864 | -91.1807
12.3868 | -91.1807
12.3871 | -91.1808
12.3874 | -91.1806
12.3878 | -91.1808
12.3881 | -91.181
12.3884 | -91.181
12.3888 | -91.181
12.3891 | -91.1809
12.3894 | -91.181
12.3897 | -91.181
12.3901 | -91.181
12.3904 | -91.1808
12.3907 | -91.1811
12.3911 | -91.1812
12.3914 | -91.1807
12.3917 | -91.1805
12.392 | -91.1808
12.3924 | -91.1808
12.3927 | -91.1808
12.393 | -91.1802
12.3934 | -91.18
12.3937 | -91.1801
12.394 | -91.18
12.3943 | -91.1797
12.3947 | -91.1798
12.395 | -91.1799
12.3953 | -91.18
12.3957 | -91.18
12.396 | -91.18
12.3963 | -91.1799
12.3967 | -91.1797
12.397 | -91.1796
12.3973 | -91.1791
12.3976 | -91.1783
12.398 | -91.1782
12.3983 | -91.1781
12.3986 | -91.1785
12.399 | -91.179
12.3993 | -91.1793
12.3996 | -91.1793
12.3999 | -91.1795
12.4003 | -91.1795
12.4006 | -91.1793
12.4009 | -91.1793
12.4013 | -91.1792
12.4016 | -91.1793
12.4019 | -91.1795
12.4022 | -91.1797
12.4026 | -91.1797
12.4029 | -91.1797
12.4032 | -91.1796
12.4036 | -91.1796
12.4039 | -91.1797
12.4042 | -91.1797
12.4046 | -91.1797
12.4049 | -91.1801
12.4052 | -91.1807
12.4055 | -91.1807
12.4059 | -91.18
12.4062 | -91.1794
12.4065 | -91.1797
12.4069 | -91.1802
12.4072 | -91.1802
12.4075 | -91.1804
12.4078 | -91.1803
12.4082 | -91.1803
12.4085 | -91.1803
12.4088 | -91.1799
12.4092 | -91.1801
12.4095 | -91.1797
12.4098 | -91.1795
12.4101 | -91.1795
12.4105 | -91.1796
12.4108 | -91.1794
12.4111 | -91.1788
12.4115 | -91.1789
12.4118 | -91.1791
12.4121 | -91.1788
12.4124 | -91.1788
12.4128 | -91.1789
12.4131 | -91.1786
12.4134 | -91.1786
12.4138 | -91.1785
12.4141 | -91.1786
12.4144 | -91.1786
12.4148 | -91.1789
12.4151 | -91.179
12.4154 | -91.179
12.4157 | -91.179
12.4161 | -91.179
12.4164 | -91.1785
12.4167 | -91.1782
12.4171 | -91.1785
12.4174 | -91.1781
12.4177 | -91.1771
12.418 | -91.176
12.4184 | -91.1755
12.4187 | -91.1752
12.419 | -91.1753
12.4194 | -91.1754
12.4197 | -91.1754
12.42 | -91.1757
12.4203 | -91.176
12.4207 | -91.176
12.421 | -91.1759
12.4213 | -91.1762
12.4217 | -91.1765
12.422 | -91.1769
12.4223 | -91.1768
12.4227 | -91.1767
12.423 | -91.1768
12.4233 | -91.1767
12.4236 | -91.1768
12.424 | -91.1774
12.4243 | -91.1779
12.4246 | -91.1781
12.425 | -91.1784
12.4253 | -91.1788
12.4256 | -91.1788
12.4259 | -91.1788
12.4263 | -91.1787
12.4266 | -91.1786
12.4269 | -91.1788
12.4273 | -91.1787
12.4276 | -91.1789
12.4279 | -91.1789
12.4282 | -91.1791
12.4286 | -91.1789
12.4289 | -91.1782
12.4292 | -91.1782
12.4296 | -91.1781
12.4299 | -91.1782
12.4302 | -91.1778
12.4306 | -91.1772
12.4309 | -91.1772
12.4312 | -91.1773
12.4315 | -91.177
12.4319 | -91.1767
12.4322 | -91.1767
12.4325 | -91.1768
12.4329 | -91.177
12.4332 | -91.177
12.4335 | -91.177
12.4338 | -91.1771
12.4342 | -91.1771
12.4345 | -91.1769
12.4348 | -91.1769
12.4352 | -91.177
12.4355 | -91.1773
12.4358 | -91.1771
12.4361 | -91.1766
12.4365 | -91.1766
12.4368 | -91.1765
12.4371 | -91.1765
12.4375 | -91.1761
12.4378 | -91.1759
12.4381 | -91.1758
12.4385 | -91.1755
12.4388 | -91.1751
12.4391 | -91.1748
12.4394 | -91.1749
12.4398 | -91.1751
12.4401 | -91.1748
12.4404 | -91.1747
12.4408 | -91.1748
12.4411 | -91.1754
12.4414 | -91.176
12.4417 | -91.176
12.4421 | -91.176
12.4424 | -91.1761
12.4427 | -91.1758
12.4431 | -91.1761
12.4434 | -91.1761
12.4437 | -91.1762
12.444 | -91.1758
12.4444 | -91.1757
12.4447 | -91.1754
12.445 | -91.175
12.4454 | -91.1747
12.4457 | -91.1744
12.446 | -91.1746
12.4464 | -91.1747
12.4467 | -91.1747
12.447 | -91.1747
12.4473 | -91.1749
12.4477 | -91.1753
12.448 | -91.175
12.4483 | -91.1748
12.4487 | -91.1749
12.449 | -91.1748
12.4493 | -91.1749
12.4496 | -91.1753
12.45 | -91.1753
12.4503 | -91.1753
12.4506 | -91.1754
12.451 | -91.1755
12.4513 | -91.176
12.4516 | -91.1763
12.4519 | -91.1765
12.4523 | -91.1762
12.4526 | -91.1762
12.4529 | -91.1762
12.4533 | -91.1757
12.4536 | -91.1756
12.4539 | -91.1758
12.4543 | -91.1759
12.4546 | -91.1755
12.4549 | -91.1748
12.4552 | -91.1744
12.4556 | -91.1744
12.4559 | -91.1745
12.4562 | -91.1742
12.4566 | -91.174
12.4569 | -91.1739
12.4572 | -91.1739
12.4575 | -91.1738
12.4579 | -91.1741
12.4582 | -91.1741
12.4585 | -91.1742
12.4589 | -91.1736
12.4592 | -91.1734
12.4595 | -91.1735
12.4598 | -91.1731
12.4602 | -91.1721
12.4605 | -91.1719
12.4608 | -91.1715
12.4612 | -91.1715
12.4615 | -91.1716
12.4618 | -91.1717
12.4621 | -91.172
12.4625 | -91.1721
12.4628 | -91.172
12.4631 | -91.1722
12.4635 | -91.1725
12.4638 | -91.1722
12.4641 | -91.1719
12.4645 | -91.1723
12.4648 | -91.1724
12.4651 | -91.1713
12.4654 | -91.1708
12.4658 | -91.1708
12.4661 | -91.1708
12.4664 | -91.1708
12.4668 | -91.1707
12.4671 | -91.1706
12.4674 | -91.1707
12.4677 | -91.1707
12.4681 | -91.1708
12.4684 | -91.1708
12.4687 | -91.171
12.4691 | -91.1712
12.4694 | -91.1714
12.4697 | -91.171
12.47 | -91.1705
12.4704 | -91.1698
12.4707 | -91.1691
12.471 | -91.1686
12.4714 | -91.1685
12.4717 | -91.1691
12.472 | -91.1691
12.4724 | -91.1682
12.4727 | -91.1671
12.473 | -91.1665
12.4733 | -91.1655
12.4737 | -91.1649
12.474 | -91.1648
12.4743 | -91.1647
12.4747 | -91.1649
12.475 | -91.165
12.4753 | -91.1649
12.4756 | -91.1642
12.476 | -91.1643
12.4763 | -91.1646
12.4766 | -91.1644
12.477 | -91.1642
12.4773 | -91.1642
12.4776 | -91.1642
12.4779 | -91.164
12.4783 | -91.1638
12.4786 | -91.1635
12.4789 | -91.1631
12.4793 | -91.1632
12.4796 | -91.1631
12.4799 | -91.1636
12.4803 | -91.1635
12.4806 | -91.1636
12.4809 | -91.1635
12.4812 | -91.1633
12.4816 | -91.1631
12.4819 | -91.1636
12.4822 | -91.1641
12.4826 | -91.1643
12.4829 | -91.1641
12.4832 | -91.1636
12.4835 | -91.1635
12.4839 | -91.1633
12.4842 | -91.1631
12.4845 | -91.1634
12.4849 | -91.163
12.4852 | -91.1629
12.4855 | -91.163
12.4858 | -91.1629
12.4862 | -91.1626
12.4865 | -91.1619
12.4868 | -91.1618
12.4872 | -91.1617
12.4875 | -91.1618
12.4878 | -91.1617
12.4882 | -91.1614
12.4885 | -91.1615
12.4888 | -91.1614
12.4891 | -91.1615
12.4895 | -91.1618
12.4898 | -91.162
12.4901 | -91.1615
12.4905 | -91.161
12.4908 | -91.1612
12.4911 | -91.161
12.4914 | -91.16
12.4918 | -91.1582
12.4921 | -91.1561
12.4924 | -91.1535
12.4928 | -91.153
12.4931 | -91.1533
12.4934 | -91.1532
12.4937 | -91.1529
12.4941 | -91.1532
12.4944 | -91.1526
12.4947 | -91.1521
12.4951 | -91.1522
12.4954 | -91.1522
12.4957 | -91.1521
12.4961 | -91.152
12.4964 | -91.1515
12.4967 | -91.1512
12.497 | -91.1507
12.4974 | -91.1505
12.4977 | -91.1503
12.498 | -91.1503
12.4984 | -91.1505
12.4987 | -91.15
12.499 | -91.1502
12.4993 | -91.1499
12.4997 | -91.1496
12.5 | -91.1497
12.5003 | -91.1495
12.5007 | -91.1488
12.501 | -91.1491
12.5013 | -91.1492
12.5016 | -91.1491
12.502 | -91.1488
12.5023 | -91.1487
12.5026 | -91.1485
12.503 | -91.1483
12.5033 | -91.1478
12.5036 | -91.1474
12.5039 | -91.1473
12.5043 | -91.1473
12.5046 | -91.1471
12.5049 | -91.1464
12.5053 | -91.1455
12.5056 | -91.1454
12.5059 | -91.1454
12.5063 | -91.1453
12.5066 | -91.1441
12.5069 | -91.1433
12.5072 | -91.1435
12.5076 | -91.1434
12.5079 | -91.1433
12.5082 | -91.1432
12.5086 | -91.1429
12.5089 | -91.1426
12.5092 | -91.1421
12.5095 | -91.1416
12.5099 | -91.1413
12.5102 | -91.1409
12.5105 | -91.1408
12.5109 | -91.1401
12.5112 | -91.1393
12.5115 | -91.1394
12.5118 | -91.1391
12.5122 | -91.1388
12.5125 | -91.1384
12.5128 | -91.1384
12.5132 | -91.1377
12.5135 | -91.1369
12.5138 | -91.1369
12.5142 | -91.1372
12.5145 | -91.1376
12.5148 | -91.1368
12.5151 | -91.1362
12.5155 | -91.136
12.5158 | -91.136
12.5161 | -91.1354
12.5165 | -91.1354
12.5168 | -91.1353
12.5171 | -91.1352
12.5174 | -91.135
12.5178 | -91.1347
12.5181 | -91.1347
12.5184 | -91.1345
12.5188 | -91.1335
12.5191 | -91.1329
12.5194 | -91.1327
12.5197 | -91.1326
12.5201 | -91.1327
12.5204 | -91.132
12.5207 | -91.1316
12.5211 | -91.1308
12.5214 | -91.1288
12.5217 | -91.1275
12.5221 | -91.1268
12.5224 | -91.1258
12.5227 | -91.1239
12.523 | -91.1211
12.5234 | -91.12
12.5237 | -91.1197
12.524 | -91.1191
12.5244 | -91.1184
12.5247 | -91.1182
12.525 | -91.1179
12.5253 | -91.1177
12.5257 | -91.1168
12.526 | -91.1164
12.5263 | -91.1159
12.5267 | -91.1156
12.527 | -91.1153
12.5273 | -91.1151
12.5276 | -91.1152
12.528 | -91.1152
12.5283 | -91.1149
12.5286 | -91.1137
12.529 | -91.1122
12.5293 | -91.1117
12.5296 | -91.1112
12.53 | -91.111
12.5303 | -91.1109
12.5306 | -91.1107
12.5309 | -91.1105
12.5313 | -91.1103
12.5316 | -91.1101
12.5319 | -91.1095
12.5323 | -91.1093
12.5326 | -91.109
12.5329 | -91.1087
12.5332 | -91.1083
12.5336 | -91.1074
12.5339 | -91.1064
12.5342 | -91.1051
12.5346 | -91.1047
12.5349 | -91.1041
12.5352 | -91.1035
12.5355 | -91.1024
12.5359 | -91.1018
12.5362 | -91.1012
12.5365 | -91.1009
12.5369 | -91.1008
12.5372 | -91.1008
12.5375 | -91.1005
12.5379 | -91.0999
12.5382 | -91.0993
12.5385 | -91.0991
12.5388 | -91.0985
12.5392 | -91.0981
12.5395 | -91.0978
12.5398 | -91.0973
12.5402 | -91.0966
12.5405 | -91.0961
12.5408 | -91.0955
12.5411 | -91.0948
12.5415 | -91.0943
12.5418 | -91.0936
12.5421 | -91.0932
12.5425 | -91.0926
12.5428 | -91.0918
12.5431 | -91.0917
12.5434 | -91.0911
12.5438 | -91.09
12.5441 | -91.0898
12.5444 | -91.0897
12.5448 | -91.0894
12.5451 | -91.0892
12.5454 | -91.089
12.5457 | -91.0886
12.5461 | -91.0877
12.5464 | -91.0861
12.5467 | -91.0856
12.5471 | -91.0852
12.5474 | -91.0846
12.5477 | -91.084
12.5481 | -91.0833
12.5484 | -91.0829
12.5487 | -91.0823
12.549 | -91.0817
12.5494 | -91.0806
12.5497 | -91.0801
12.55 | -91.0799
12.5504 | -91.0794
12.5507 | -91.0786
12.551 | -91.0784
12.5513 | -91.0782
12.5517 | -91.0776
12.552 | -91.0775
12.5523 | -91.0774
12.5527 | -91.0769
12.553 | -91.0764
12.5533 | -91.0757
12.5536 | -91.0751
12.554 | -91.0741
12.5543 | -91.0737
12.5546 | -91.0734
12.555 | -91.073
12.5553 | -91.0732
12.5556 | -91.0728
12.556 | -91.0724
12.5563 | -91.072
12.5566 | -91.0716
12.5569 | -91.0712
12.5573 | -91.0709
12.5576 | -91.07
12.5579 | -91.0683
12.5583 | -91.0672
12.5586 | -91.0669
12.5589 | -91.0661
12.5592 | -91.0656
12.5596 | -91.0652
12.5599 | -91.0646
12.5602 | -91.0637
12.5606 | -91.0622
12.5609 | -91.0614
12.5612 | -91.061
12.5615 | -91.0605
12.5619 | -91.0593
12.5622 | -91.0589
12.5625 | -91.0584
12.5629 | -91.058
12.5632 | -91.0576
12.5635 | -91.0564
12.5639 | -91.0558
12.5642 | -91.0553
12.5645 | -91.0545
12.5648 | -91.0542
12.5652 | -91.0525
12.5655 | -91.0511
12.5658 | -91.0506
12.5662 | -91.0503
12.5665 | -91.0501
12.5668 | -91.0492
12.5671 | -91.0482
12.5675 | -91.0471
12.5678 | -91.0467
12.5681 | -91.046
12.5685 | -91.0455
12.5688 | -91.0435
12.5691 | -91.0421
12.5694 | -91.0416
12.5698 | -91.0412
12.5701 | -91.0411
12.5704 | -91.0404
12.5708 | -91.0395
12.5711 | -91.0383
12.5714 | -91.0375
12.5718 | -91.0369
12.5721 | -91.0347
12.5724 | -91.0323
12.5727 | -91.0319
12.5731 | -91.0294
12.5734 | -91.0275
12.5737 | -91.0253
12.5741 | -91.0242
12.5744 | -91.0235
12.5747 | -91.0228
12.575 | -91.0218
12.5754 | -91.0204
12.5757 | -91.0196
12.576 | -91.0186
12.5764 | -91.0181
12.5767 | -91.0174
12.577 | -91.0168
12.5773 | -91.0159
12.5777 | -91.0153
12.578 | -91.0146
12.5783 | -91.0141
12.5787 | -91.0131
12.579 | -91.0124
12.5793 | -91.0115
12.5797 | -91.0109
12.58 | -91.0102
12.5803 | -91.0097
12.5806 | -91.0089
12.581 | -91.0085
12.5813 | -91.008
12.5816 | -91.0071
12.582 | -91.0065
12.5823 | -91.0058
12.5826 | -91.0045
12.5829 | -91.0037
12.5833 | -91.0034
12.5836 | -91.0029
12.5839 | -91.0022
12.5843 | -91.0015
12.5846 | -91.0009
12.5849 | -91.0002
12.5852 | -90.9995
12.5856 | -90.9988
12.5859 | -90.9981
12.5862 | -90.9975
12.5866 | -90.9968
12.5869 | -90.9957
12.5872 | -90.9945
12.5876 | -90.9938
12.5879 | -90.9929
12.5882 | -90.9914
12.5885 | -90.9905
12.5889 | -90.9887
12.5892 | -90.9867
12.5895 | -90.9858
12.5899 | -90.9852
12.5902 | -90.9842
12.5905 | -90.9833
12.5908 | -90.9824
12.5912 | -90.9813
12.5915 | -90.9805
12.5918 | -90.9799
12.5922 | -90.9787
12.5925 | -90.9779
12.5928 | -90.9771
12.5931 | -90.9764
12.5935 | -90.9763
12.5938 | -90.9758
12.5941 | -90.9752
12.5945 | -90.9741
12.5948 | -90.9732
12.5951 | -90.9726
12.5954 | -90.9716
12.5958 | -90.9701
12.5961 | -90.9679
12.5964 | -90.9663
12.5968 | -90.9651
12.5971 | -90.9638
12.5974 | -90.9624
12.5978 | -90.9617
12.5981 | -90.9604
12.5984 | -90.9586
12.5987 | -90.9565
12.5991 | -90.9539
12.5994 | -90.9517
12.5997 | -90.95
12.6001 | -90.9489
12.6004 | -90.9481
12.6007 | -90.9468
12.601 | -90.9457
12.6014 | -90.9447
12.6017 | -90.9429
12.602 | -90.942
12.6024 | -90.9412
12.6027 | -90.9404
12.603 | -90.9395
12.6033 | -90.9388
12.6037 | -90.9379
12.604 | -90.9369
12.6043 | -90.9359
12.6047 | -90.9351
12.605 | -90.9343
12.6053 | -90.933
12.6057 | -90.9322
12.606 | -90.9317
12.6063 | -90.9311
12.6066 | -90.9304
12.607 | -90.9296
12.6073 | -90.9287
12.6076 | -90.9272
12.608 | -90.9256
12.6083 | -90.9246
12.6086 | -90.9232
12.6089 | -90.9195
12.6093 | -90.9167
12.6096 | -90.9164
12.6099 | -90.9152
12.6103 | -90.9139
12.6106 | -90.9133
12.6109 | -90.9119
12.6112 | -90.9098
12.6116 | -90.9091
12.6119 | -90.9086
12.6122 | -90.9075
12.6126 | -90.9064
12.6129 | -90.9052
12.6132 | -90.9032
12.6136 | -90.9018
12.6139 | -90.9005
12.6142 | -90.8988
12.6145 | -90.8973
12.6149 | -90.8959
12.6152 | -90.8949
12.6155 | -90.894
12.6159 | -90.893
12.6162 | -90.8924
12.6165 | -90.8916
12.6168 | -90.8901
12.6172 | -90.8877
12.6175 | -90.8855
12.6178 | -90.8841
12.6182 | -90.8829
12.6185 | -90.8816
12.6188 | -90.8794
12.6191 | -90.8781
12.6195 | -90.876
12.6198 | -90.875
12.6201 | -90.8732
12.6205 | -90.8707
12.6208 | -90.8697
12.6211 | -90.8686
12.6215 | -90.8676
12.6218 | -90.8664
12.6221 | -90.8646
12.6224 | -90.8634
12.6228 | -90.8621
12.6231 | -90.8611
12.6234 | -90.8602
12.6238 | -90.8593
12.6241 | -90.8585
12.6244 | -90.8574
12.6247 | -90.8566
12.6251 | -90.8557
12.6254 | -90.8547
12.6257 | -90.8531
12.6261 | -90.8519
12.6264 | -90.8506
12.6267 | -90.8484
12.627 | -90.8466
12.6274 | -90.8443
12.6277 | -90.8425
12.628 | -90.8415
12.6284 | -90.8409
12.6287 | -90.8401
12.629 | -90.8389
12.6294 | -90.8379
12.6297 | -90.8361
12.63 | -90.832
12.6303 | -90.8304
12.6307 | -90.8293
12.631 | -90.8275
12.6313 | -90.8255
12.6317 | -90.8244
12.632 | -90.8232
12.6323 | -90.8217
12.6326 | -90.8202
12.633 | -90.8191
12.6333 | -90.818
12.6336 | -90.8169
12.634 | -90.8154
12.6343 | -90.8141
12.6346 | -90.8129
12.6349 | -90.8114
12.6353 | -90.8101
12.6356 | -90.8092
12.6359 | -90.8083
12.6363 | -90.8061
12.6366 | -90.8029
12.6369 | -90.8005
12.6373 | -90.7986
12.6376 | -90.7974
12.6379 | -90.7962
12.6382 | -90.7947
12.6386 | -90.7926
12.6389 | -90.7918
12.6392 | -90.7904
12.6396 | -90.7869
12.6399 | -90.7842
12.6402 | -90.7828
12.6405 | -90.7817
12.6409 | -90.7806
12.6412 | -90.7796
12.6415 | -90.7784
12.6419 | -90.7768
12.6422 | -90.7752
12.6425 | -90.7738
12.6428 | -90.7726
12.6432 | -90.7711
12.6435 | -90.7694
12.6438 | -90.7678
12.6442 | -90.7662
12.6445 | -90.7647
12.6448 | -90.7631
12.6451 | -90.7618
12.6455 | -90.7602
12.6458 | -90.7583
12.6461 | -90.757
12.6465 | -90.756
12.6468 | -90.7546
12.6471 | -90.7527
12.6475 | -90.7513
12.6478 | -90.7496
12.6481 | -90.7484
12.6484 | -90.747
12.6488 | -90.7453
12.6491 | -90.7431
12.6494 | -90.742
12.6498 | -90.741
12.6501 | -90.7391
12.6504 | -90.737
12.6507 | -90.7351
12.6511 | -90.7338
12.6514 | -90.733
12.6517 | -90.7311
12.6521 | -90.7274
12.6524 | -90.7231
12.6527 | -90.7214
12.653 | -90.7204
12.6534 | -90.7192
12.6537 | -90.7175
12.654 | -90.7162
12.6544 | -90.715
12.6547 | -90.7135
12.655 | -90.7121
12.6554 | -90.7104
12.6557 | -90.709
12.656 | -90.7071
12.6563 | -90.7058
12.6567 | -90.7043
12.657 | -90.702
12.6573 | -90.7005
12.6577 | -90.6987
12.658 | -90.6961
12.6583 | -90.6944
12.6586 | -90.6909
12.659 | -90.6893
12.6593 | -90.688
12.6596 | -90.6869
12.66 | -90.6859
12.6603 | -90.6837
12.6606 | -90.6811
12.6609 | -90.6798
12.6613 | -90.6785
12.6616 | -90.6754
12.6619 | -90.6714
12.6623 | -90.6699
12.6626 | -90.6676
12.6629 | -90.6654
12.6633 | -90.6639
12.6636 | -90.6623
12.6639 | -90.6595
12.6642 | -90.6571
12.6646 | -90.6557
12.6649 | -90.6535
12.6652 | -90.652
12.6656 | -90.6503
12.6659 | -90.6488
12.6662 | -90.6463
12.6665 | -90.644
12.6669 | -90.6425
12.6672 | -90.6406
12.6675 | -90.639
12.6679 | -90.6376
12.6682 | -90.6354
12.6685 | -90.6336
12.6688 | -90.6311
12.6692 | -90.6269
12.6695 | -90.6251
12.6698 | -90.6236
12.6702 | -90.6216
12.6705 | -90.6196
12.6708 | -90.6162
12.6712 | -90.6142
12.6715 | -90.6123
12.6718 | -90.6106
12.6721 | -90.6088
12.6725 | -90.6071
12.6728 | -90.6057
12.6731 | -90.6042
12.6735 | -90.602
12.6738 | -90.5999
12.6741 | -90.5982
12.6744 | -90.5964
12.6748 | -90.5945
12.6751 | -90.5922
12.6754 | -90.5897
12.6758 | -90.5872
12.6761 | -90.5856
12.6764 | -90.5843
12.6767 | -90.5827
12.6771 | -90.581
12.6774 | -90.5793
12.6777 | -90.5763
12.6781 | -90.574
12.6784 | -90.5723
12.6787 | -90.57
12.6791 | -90.5672
12.6794 | -90.5657
12.6797 | -90.5636
12.68 | -90.5617
12.6804 | -90.5605
12.6807 | -90.5588
12.681 | -90.5569
12.6814 | -90.5551
12.6817 | -90.5519
12.682 | -90.5481
12.6823 | -90.5458
12.6827 | -90.5441
12.683 | -90.5422
12.6833 | -90.54
12.6837 | -90.5371
12.684 | -90.5344
12.6843 | -90.5327
12.6846 | -90.5308
12.685 | -90.5278
12.6853 | -90.525
12.6856 | -90.5232
12.686 | -90.5209
12.6863 | -90.5186
12.6866 | -90.5166
12.6869 | -90.5141
12.6873 | -90.5117
12.6876 | -90.5086
12.6879 | -90.5061
12.6883 | -90.5042
12.6886 | -90.502
12.6889 | -90.5001
12.6893 | -90.4974
12.6896 | -90.4942
12.6899 | -90.4924
12.6902 | -90.4912
12.6906 | -90.4895
12.6909 | -90.4867
12.6912 | -90.4841
12.6916 | -90.4822
12.6919 | -90.48
12.6922 | -90.4784
12.6925 | -90.4765
12.6929 | -90.4749
12.6932 | -90.4732
12.6935 | -90.4707
12.6939 | -90.467
12.6942 | -90.4638
12.6945 | -90.4617
12.6949 | -90.4594
12.6952 | -90.4578
12.6955 | -90.4558
12.6958 | -90.4537
12.6962 | -90.4518
12.6965 | -90.4486
12.6968 | -90.4467
12.6972 | -90.4447
12.6975 | -90.4422
12.6978 | -90.4397
12.6981 | -90.437
12.6985 | -90.4338
12.6988 | -90.4311
12.6991 | -90.4281
12.6995 | -90.4258
12.6998 | -90.4235
12.7001 | -90.4212
12.7004 | -90.4193
12.7008 | -90.4179
12.7011 | -90.4159
12.7014 | -90.4139
12.7018 | -90.4117
12.7021 | -90.4095
12.7024 | -90.4075
12.7027 | -90.4052
12.7031 | -90.4031
12.7034 | -90.4009
12.7037 | -90.398
12.7041 | -90.3952
12.7044 | -90.3925
12.7047 | -90.3902
12.7051 | -90.388
12.7054 | -90.386
12.7057 | -90.3836
12.706 | -90.3801
12.7064 | -90.3781
12.7067 | -90.3753
12.707 | -90.3727
12.7074 | -90.3705
12.7077 | -90.3689
12.708 | -90.3668
12.7083 | -90.3645
12.7087 | -90.3623
12.709 | -90.3602
12.7093 | -90.3577
12.7097 | -90.3547
12.71 | -90.352
12.7103 | -90.3501
12.7106 | -90.3476
12.711 | -90.3453
12.7113 | -90.3427
12.7116 | -90.3398
12.712 | -90.3373
12.7123 | -90.3355
12.7126 | -90.3339
12.713 | -90.3317
12.7133 | -90.3295
12.7136 | -90.3273
12.7139 | -90.3252
12.7143 | -90.3226
12.7146 | -90.3198
12.7149 | -90.3177
12.7153 | -90.3163
12.7156 | -90.3135
12.7159 | -90.3116
12.7162 | -90.3096
12.7166 | -90.3077
12.7169 | -90.3051
12.7172 | -90.3025
12.7176 | -90.2999
12.7179 | -90.2972
12.7182 | -90.2941
12.7185 | -90.2917
12.7189 | -90.289
12.7192 | -90.2867
12.7195 | -90.2848
12.7199 | -90.2826
12.7202 | -90.2801
12.7205 | -90.2777
12.7209 | -90.2756
12.7212 | -90.2733
12.7215 | -90.2705
12.7218 | -90.268
12.7222 | -90.266
12.7225 | -90.2637
12.7228 | -90.2616
12.7232 | -90.2595
12.7235 | -90.257
12.7238 | -90.2546
12.7241 | -90.2522
12.7245 | -90.2494
12.7248 | -90.2471
12.7251 | -90.244
12.7255 | -90.2415
12.7258 | -90.239
12.7261 | -90.2369
12.7264 | -90.2345
12.7268 | -90.2321
12.7271 | -90.2298
12.7274 | -90.2278
12.7278 | -90.2252
12.7281 | -90.2228
12.7284 | -90.2208
12.7288 | -90.2186
12.7291 | -90.2162
12.7294 | -90.2135
12.7297 | -90.2106
12.7301 | -90.2079
12.7304 | -90.2055
12.7307 | -90.2028
12.7311 | -90.2005
12.7314 | -90.1982
12.7317 | -90.1955
12.732 | -90.192
12.7324 | -90.1889
12.7327 | -90.1862
12.733 | -90.184
12.7334 | -90.182
12.7337 | -90.1798
12.734 | -90.1772
12.7343 | -90.1743
12.7347 | -90.1705
12.735 | -90.168
12.7353 | -90.1656
12.7357 | -90.1633
12.736 | -90.1607
12.7363 | -90.1581
12.7366 | -90.1549
12.737 | -90.1526
12.7373 | -90.1502
12.7376 | -90.1481
12.738 | -90.1457
12.7383 | -90.1434
12.7386 | -90.1412
12.739 | -90.1389
12.7393 | -90.1363
12.7396 | -90.1337
12.7399 | -90.1313
12.7403 | -90.1286
12.7406 | -90.1253
12.7409 | -90.123
12.7413 | -90.1202
12.7416 | -90.1159
12.7419 | -90.1115
12.7422 | -90.1075
12.7426 | -90.104
12.7429 | -90.1012
12.7432 | -90.0985
12.7436 | -90.0958
12.7439 | -90.0927
12.7442 | -90.0905
12.7445 | -90.0884
12.7449 | -90.086
12.7452 | -90.0831
12.7455 | -90.0806
12.7459 | -90.0781
12.7462 | -90.0757
12.7465 | -90.0733
12.7469 | -90.0712
12.7472 | -90.0687
12.7475 | -90.0662
12.7478 | -90.0636
12.7482 | -90.0613
12.7485 | -90.0583
12.7488 | -90.0558
12.7492 | -90.0537
12.7495 | -90.0512
12.7498 | -90.0489
12.7501 | -90.0466
12.7505 | -90.0441
12.7508 | -90.0415
12.7511 | -90.039
12.7515 | -90.0363
12.7518 | -90.0332
12.7521 | -90.0304
12.7524 | -90.0279
12.7528 | -90.0249
12.7531 | -90.0222
12.7534 | -90.0191
12.7538 | -90.0169
12.7541 | -90.0147
12.7544 | -90.0119
12.7548 | -90.009
12.7551 | -90.006
12.7554 | -90.0034
12.7557 | -90.0012
12.7561 | -89.9989
12.7564 | -89.9965
12.7567 | -89.994
12.7571 | -89.9916
12.7574 | -89.9892
12.7577 | -89.9868
12.758 | -89.9838
12.7584 | -89.9805
12.7587 | -89.9787
12.759 | -89.977
12.7594 | -89.9746
12.7597 | -89.9716
12.76 | -89.969
12.7603 | -89.9666
12.7607 | -89.9639
12.761 | -89.9613
12.7613 | -89.9591
12.7617 | -89.9567
12.762 | -89.9541
12.7623 | -89.9516
12.7627 | -89.9482
12.763 | -89.9435
12.7633 | -89.9408
12.7636 | -89.9381
12.764 | -89.9356
12.7643 | -89.9329
12.7646 | -89.9302
12.765 | -89.9274
12.7653 | -89.9249
12.7656 | -89.9225
12.7659 | -89.9199
12.7663 | -89.9164
12.7666 | -89.9128
12.7669 | -89.9099
12.7673 | -89.9074
12.7676 | -89.9052
12.7679 | -89.9025
12.7682 | -89.8997
12.7686 | -89.8971
12.7689 | -89.8945
12.7692 | -89.8915
12.7696 | -89.8887
12.7699 | -89.8858
12.7702 | -89.8828
12.7706 | -89.88
12.7709 | -89.8774
12.7712 | -89.8751
12.7715 | -89.8729
12.7719 | -89.87
12.7722 | -89.8673
12.7725 | -89.8637
12.7729 | -89.8611
12.7732 | -89.8585
12.7735 | -89.8559
12.7738 | -89.8532
12.7742 | -89.8507
12.7745 | -89.8479
12.7748 | -89.8452
12.7752 | -89.8425
12.7755 | -89.8399
12.7758 | -89.8372
12.7761 | -89.8339
12.7765 | -89.8306
12.7768 | -89.8275
12.7771 | -89.8246
12.7775 | -89.8218
12.7778 | -89.8192
12.7781 | -89.8166
12.7785 | -89.814
12.7788 | -89.8114
12.7791 | -89.8095
12.7794 | -89.8074
12.7798 | -89.8044
12.7801 | -89.8012
12.7804 | -89.7984
12.7808 | -89.7958
12.7811 | -89.793
12.7814 | -89.7904
12.7817 | -89.7876
12.7821 | -89.784
12.7824 | -89.7808
12.7827 | -89.7782
12.7831 | -89.7755
12.7834 | -89.773
12.7837 | -89.7703
12.784 | -89.7676
12.7844 | -89.7649
12.7847 | -89.7621
12.785 | -89.7589
12.7854 | -89.7561
12.7857 | -89.7534
12.786 | -89.7503
12.7864 | -89.7477
12.7867 | -89.745
12.787 | -89.742
12.7873 | -89.739
12.7877 | -89.7361
12.788 | -89.7334
12.7883 | -89.7307
12.7887 | -89.7279
12.789 | -89.725
12.7893 | -89.722
12.7896 | -89.7192
12.79 | -89.7165
12.7903 | -89.7137
12.7906 | -89.7108
12.791 | -89.7079
12.7913 | -89.7056
12.7916 | -89.7017
12.7919 | -89.6991
12.7923 | -89.6959
12.7926 | -89.6931
12.7929 | -89.6904
12.7933 | -89.6877
12.7936 | -89.6849
12.7939 | -89.6823
12.7942 | -89.6796
12.7946 | -89.6768
12.7949 | -89.6737
12.7952 | -89.6708
12.7956 | -89.6678
12.7959 | -89.665
12.7962 | -89.6623
12.7966 | -89.6597
12.7969 | -89.6568
12.7972 | -89.6539
12.7975 | -89.651
12.7979 | -89.6481
12.7982 | -89.6452
12.7985 | -89.6424
12.7989 | -89.6398
12.7992 | -89.6374
12.7995 | -89.6347
12.7998 | -89.632
12.8002 | -89.6296
12.8005 | -89.6266
12.8008 | -89.6235
12.8012 | -89.6197
12.8015 | -89.6169
12.8018 | -89.6142
12.8021 | -89.6115
12.8025 | -89.6086
12.8028 | -89.6057
12.8031 | -89.6029
12.8035 | -89.6002
12.8038 | -89.5975
12.8041 | -89.595
12.8045 | -89.5924
12.8048 | -89.5897
12.8051 | -89.5867
12.8054 | -89.5837
12.8058 | -89.581
12.8061 | -89.5786
12.8064 | -89.5757
12.8068 | -89.573
12.8071 | -89.5696
12.8074 | -89.5663
12.8077 | -89.5634
12.8081 | -89.5606
12.8084 | -89.5578
12.8087 | -89.5547
12.8091 | -89.551
12.8094 | -89.5482
12.8097 | -89.5458
12.81 | -89.5432
12.8104 | -89.5391
12.8107 | -89.5359
12.811 | -89.5328
12.8114 | -89.5299
12.8117 | -89.5273
12.812 | -89.5246
12.8124 | -89.5219
12.8127 | -89.5189
12.813 | -89.516
12.8133 | -89.5134
12.8137 | -89.5108
12.814 | -89.5077
12.8143 | -89.5044
12.8147 | -89.5016
12.815 | -89.4986
12.8153 | -89.4954
12.8156 | -89.4923
12.816 | -89.4895
12.8163 | -89.4868
12.8166 | -89.484
12.817 | -89.4808
12.8173 | -89.4779
12.8176 | -89.4752
12.8179 | -89.4722
12.8183 | -89.4692
12.8186 | -89.4665
12.8189 | -89.4639
12.8193 | -89.4611
12.8196 | -89.4581
12.8199 | -89.4551
12.8203 | -89.4523
12.8206 | -89.449
12.8209 | -89.4463
12.8212 | -89.4433
12.8216 | -89.4402
12.8219 | -89.4374
12.8222 | -89.4347
12.8226 | -89.432
12.8229 | -89.4293
12.8232 | -89.4266
12.8235 | -89.4239
12.8239 | -89.4213
12.8242 | -89.4185
12.8245 | -89.4159
12.8249 | -89.4132
12.8252 | -89.4104
12.8255 | -89.4076
12.8258 | -89.4046
12.8262 | -89.402
12.8265 | -89.3993
12.8268 | -89.3962
12.8272 | -89.3934
12.8275 | -89.3908
12.8278 | -89.3879
12.8282 | -89.3851
12.8285 | -89.3824
12.8288 | -89.3796
12.8291 | -89.3769
12.8295 | -89.374
12.8298 | -89.3713
12.8301 | -89.3684
12.8305 | -89.3655
12.8308 | -89.3625
12.8311 | -89.3594
12.8314 | -89.3561
12.8318 | -89.3531
12.8321 | -89.3505
12.8324 | -89.3478
12.8328 | -89.3447
12.8331 | -89.3417
12.8334 | -89.3389
12.8337 | -89.3361
12.8341 | -89.3333
12.8344 | -89.3303
12.8347 | -89.3271
12.8351 | -89.3238
12.8354 | -89.3207
12.8357 | -89.3174
12.8361 | -89.3147
12.8364 | -89.3118
12.8367 | -89.3087
12.837 | -89.3055
12.8374 | -89.3027
12.8377 | -89.3002
12.838 | -89.2977
12.8384 | -89.295
12.8387 | -89.2923
12.839 | -89.2895
12.8393 | -89.2869
12.8397 | -89.2842
12.84 | -89.281
12.8403 | -89.2779
12.8407 | -89.2751
12.841 | -89.2724
12.8413 | -89.2696
12.8416 | -89.2665
12.842 | -89.2637
12.8423 | -89.261
12.8426 | -89.2582
12.843 | -89.2554
12.8433 | -89.2526
12.8436 | -89.2499
12.8439 | -89.2473
12.8443 | -89.2446
12.8446 | -89.2419
12.8449 | -89.2393
12.8453 | -89.2364
12.8456 | -89.2337
12.8459 | -89.231
12.8463 | -89.2283
12.8466 | -89.2255
12.8469 | -89.2227
12.8472 | -89.22
12.8476 | -89.2172
12.8479 | -89.2142
12.8482 | -89.2115
12.8486 | -89.2086
12.8489 | -89.2058
12.8492 | -89.2031
12.8495 | -89.2003
12.8499 | -89.1974
12.8502 | -89.1948
12.8505 | -89.1925
12.8509 | -89.1901
12.8512 | -89.1873
12.8515 | -89.1846
12.8518 | -89.1821
12.8522 | -89.1793
12.8525 | -89.1764
12.8528 | -89.1736
12.8532 | -89.1704
12.8535 | -89.1673
12.8538 | -89.1647
12.8542 | -89.162
12.8545 | -89.1592
12.8548 | -89.1566
12.8551 | -89.1537
12.8555 | -89.1508
12.8558 | -89.148
12.8561 | -89.1452
12.8565 | -89.1426
12.8568 | -89.1399
12.8571 | -89.1371
12.8574 | -89.1346
12.8578 | -89.1321
12.8581 | -89.1293
12.8584 | -89.1266
12.8588 | -89.1238
12.8591 | -89.1214
12.8594 | -89.1187
12.8597 | -89.116
12.8601 | -89.1133
12.8604 | -89.1106
12.8607 | -89.1079
12.8611 | -89.1052
12.8614 | -89.1028
12.8617 | -89.1
12.8621 | -89.0973
12.8624 | -89.0946
12.8627 | -89.092
12.863 | -89.0894
12.8634 | -89.0868
12.8637 | -89.0842
12.864 | -89.0814
12.8644 | -89.0784
12.8647 | -89.0758
12.865 | -89.0731
12.8653 | -89.0702
12.8657 | -89.0674
12.866 | -89.0647
12.8663 | -89.062
12.8667 | -89.0594
12.867 | -89.0567
12.8673 | -89.0541
12.8676 | -89.0516
12.868 | -89.049
12.8683 | -89.0463
12.8686 | -89.0438
12.869 | -89.0412
12.8693 | -89.0384
12.8696 | -89.0358
12.87 | -89.0331
12.8703 | -89.0303
12.8706 | -89.0276
12.8709 | -89.025
12.8713 | -89.0223
12.8716 | -89.0195
12.8719 | -89.0168
12.8723 | -89.0139
12.8726 | -89.0113
12.8729 | -89.0087
12.8732 | -89.0061
12.8736 | -89.0035
12.8739 | -89.001
12.8742 | -88.9983
12.8746 | -88.9957
12.8749 | -88.9931
12.8752 | -88.9907
12.8755 | -88.9882
12.8759 | -88.9857
12.8762 | -88.9832
12.8765 | -88.9806
12.8769 | -88.9783
12.8772 | -88.9757
12.8775 | -88.9728
12.8779 | -88.9699
12.8782 | -88.9673
12.8785 | -88.965
12.8788 | -88.9626
12.8792 | -88.9598
12.8795 | -88.957
12.8798 | -88.9546
12.8802 | -88.9521
12.8805 | -88.9497
12.8808 | -88.9473
12.8811 | -88.9447
12.8815 | -88.9422
12.8818 | -88.9397
12.8821 | -88.9371
12.8825 | -88.9345
12.8828 | -88.9319
12.8831 | -88.9294
12.8834 | -88.927
12.8838 | -88.9245
12.8841 | -88.922
12.8844 | -88.9192
12.8848 | -88.9167
12.8851 | -88.9141
12.8854 | -88.9118
12.8857 | -88.9095
12.8861 | -88.9071
12.8864 | -88.9045
12.8867 | -88.9021
12.8871 | -88.8996
12.8874 | -88.897
12.8877 | -88.8946
12.8881 | -88.8919
12.8884 | -88.8892
12.8887 | -88.8867
12.889 | -88.8841
12.8894 | -88.8816
12.8897 | -88.8792
12.89 | -88.8766
12.8904 | -88.8745
12.8907 | -88.872
12.891 | -88.8695
12.8913 | -88.8669
12.8917 | -88.8645
12.892 | -88.8621
12.8923 | -88.8598
12.8927 | -88.8571
12.893 | -88.8547
12.8933 | -88.8524
12.8936 | -88.8502
12.894 | -88.8478
12.8943 | -88.8453
12.8946 | -88.8428
12.895 | -88.8403
12.8953 | -88.8378
12.8956 | -88.8352
12.896 | -88.8329
12.8963 | -88.8304
12.8966 | -88.828
12.8969 | -88.8258
12.8973 | -88.8234
12.8976 | -88.8208
12.8979 | -88.8184
12.8983 | -88.8158
12.8986 | -88.8133
12.8989 | -88.811
12.8992 | -88.8086
12.8996 | -88.8062
12.8999 | -88.8037
12.9002 | -88.8014
12.9006 | -88.7994
12.9009 | -88.797
12.9012 | -88.7946
12.9015 | -88.7922
12.9019 | -88.7897
12.9022 | -88.7872
12.9025 | -88.7847
12.9029 | -88.7823
12.9032 | -88.7798
12.9035 | -88.7775
12.9039 | -88.7752
12.9042 | -88.7729
12.9045 | -88.7704
12.9048 | -88.7676
12.9052 | -88.7646
12.9055 | -88.7621
12.9058 | -88.7597
12.9062 | -88.7573
12.9065 | -88.7549
12.9068 | -88.7527
12.9071 | -88.7504
12.9075 | -88.7478
12.9078 | -88.7451
12.9081 | -88.7425
12.9085 | -88.7401
12.9088 | -88.7377
12.9091 | -88.7354
12.9094 | -88.7331
12.9098 | -88.7308
12.9101 | -88.7285
12.9104 | -88.7262
12.9108 | -88.7239
12.9111 | -88.7215
12.9114 | -88.7195
12.9118 | -88.7173
12.9121 | -88.7149
12.9124 | -88.7126
12.9127 | -88.7102
12.9131 | -88.7078
12.9134 | -88.7055
12.9137 | -88.7032
12.9141 | -88.701
12.9144 | -88.6985
12.9147 | -88.696
12.915 | -88.6941
12.9154 | -88.6919
12.9157 | -88.6898
12.916 | -88.6875
12.9164 | -88.685
12.9167 | -88.6826
12.917 | -88.6803
12.9173 | -88.678
12.9177 | -88.6758
12.918 | -88.6738
12.9183 | -88.6716
12.9187 | -88.6693
12.919 | -88.6671
12.9193 | -88.6646
12.9197 | -88.662
12.92 | -88.6597
12.9203 | -88.6575
12.9206 | -88.6554
12.921 | -88.6531
12.9213 | -88.6508
12.9216 | -88.6489
12.922 | -88.6467
12.9223 | -88.6445
12.9226 | -88.6423
12.9229 | -88.64
12.9233 | -88.6377
12.9236 | -88.6355
12.9239 | -88.6333
12.9243 | -88.631
12.9246 | -88.6293
12.9249 | -88.627
12.9252 | -88.6251
12.9256 | -88.6228
12.9259 | -88.6208
12.9262 | -88.6189
12.9266 | -88.6167
12.9269 | -88.6147
12.9272 | -88.6127
12.9276 | -88.6105
12.9279 | -88.6082
12.9282 | -88.606
12.9285 | -88.6039
12.9289 | -88.6019
12.9292 | -88.5999
12.9295 | -88.5978
12.9299 | -88.5956
12.9302 | -88.5936
12.9305 | -88.5914
12.9308 | -88.5891
12.9312 | -88.587
12.9315 | -88.5849
12.9318 | -88.5828
12.9322 | -88.5808
12.9325 | -88.5787
12.9328 | -88.5766
12.9331 | -88.5743
12.9335 | -88.572
12.9338 | -88.5699
12.9341 | -88.5683
12.9345 | -88.5661
12.9348 | -88.5639
12.9351 | -88.5618
12.9354 | -88.5596
12.9358 | -88.5575
12.9361 | -88.5552
12.9364 | -88.5527
12.9368 | -88.5504
12.9371 | -88.5483
12.9374 | -88.5463
12.9378 | -88.5442
12.9381 | -88.5422
12.9384 | -88.5402
12.9387 | -88.5383
12.9391 | -88.5363
12.9394 | -88.5344
12.9397 | -88.5324
12.9401 | -88.5302
12.9404 | -88.5281
12.9407 | -88.5262
12.941 | -88.5243
12.9414 | -88.5226
12.9417 | -88.5206
12.942 | -88.5189