## CALIBRATION_CERTIFICATE
SENSOR IDENTIFICATION |  |  |  | CUSTOMER
Sensor Serial Number | MB3a | 73 |  | CTBTO
Transducer Block | 02-90C/P | 71 |  | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
Hood | 02-92B/P | 60
ENVIRONNEMENTAL PARAMETERS |  |  |  | PROCEDURE REFERENCE
Atmospheric pressure | 1013.7 | hPa |  | PR 000 150_-_CAL
Humidity | 19.2 | %/Hr |  | Measurement device
Internal Temperature of Calibration Bench | 23.1 | °C |  | IS CALIBRATOR
AUX Sensor pressure | 1017 | hPa |  | Sensor SoH at test start
AUX Sensor Temperature | 23.25 | °C |  | OK
SENSITIVITY AT 1 Hz
Output Parameters | Theoretical Value | Measured Value @1Hz | Margin (±) | Unit | Result
Sensitivity of Pressure Derivative output | 2 | 2 | 0.04 | mV/Pa/sec | OK | Mode d'emploi pour créer un constat de vérification
Sensitivity of Pressure output | 20 | 20 | 0.4 | mV/Pa | OK | Cliquer sur "DEMARRER LA SAISIE"
Sensitivity of Calibration coil | 6 | 5.96 | 0.13 | Pa/V | OK | Suivre les instructions affichées
Une fois la partie SENSOR IDENTIFICATION remplie, renseigner la partie ENVIRONNEMENTAL PARAMETERS (les données environnementales)
TRANSFER FUNCTION | [5 empty] | Renseigner la partie SENSITIVITY AT 1Hz :les sensibilités en voie de dérivée de Pression (Pressure Derivative), en voie de pression (Pressure) et de Calibration (calibration Coil)
Electrical |  | Acoustical |  |  |  | Une sauvegarde temporaire est toujours possible si le fichier doit être modifié (Fichier "enregistrer")
Vérifier toutes les données saisie notamment la partie CUSTOMER ("Client") et la partie VISA (Noms des personnes et date d'ajustage)
Cliquer sur "ENREGISTREMENT DEFINITIF" (Attention le fichier sera enregistré et verrouillé dans le dossier courant, il faut retirer manuellement la protection)
SENSOR SELF NOISE |  |  |  | Decision : considering the specifications given by the manufacturer, this sensor is
CONFORM
VISA
|  |  |  | Adjusted by | A.HUE
|  |  |  | on | 2016-03-15
|  |  |  | Approved by | A.HUE
|  |  |  | Qualification | Optronic Engineer
Self Noise at 1 HZ | 8.3657e-05 | Pa/√Hz |  | on | 2016-03-17 | -81.55 | (dBPa/√Hz)
| < 0,13 | mPa/√Hz
OK
The reproduction of this verification report is only authorized in the form of the full photographic facsimile. This document cannot be used instead of a calibration certificate; this document is established following the recommendations of X 07-011 which define the verification report. It can be used to demonstrate the link between the measure equipment and the standard national or international equipment, provided it meets the recommendations of the documentation booklet X 07-015
[8 empty] | Client | Adresse Client
[8 empty] | CEA | C.E.A. Dam. Bruyères Domaine du Grand Rué 91297 ARPAJON CEDEX FRANCE
[8 empty] | CTBTO | Vienna International Centre Wagramer strasse.,5 A-1400 VIENNA, AUSTRIA
[8 empty] | IPGP | Université Paris Diderot Paris 7 Bat Lamarck UFR STEP 35 rue H.Brion 6ème étage 75205 Paris CEDEX 13
[8 empty] | SANDIA | C/O SANDIATIONAL LABS ALBUQUERQUE, NM 87123 UNITED STATES
[8 empty] | ISLA | Haiwaï Institute of Geophysics & Planetology 73-984 Makako Bay Drive KAILUA-KONA HAWAÏ
ENREGISTREMENT DEFINITIF
DEMARRER LA SAISIE

## ELEC_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.2297 | 12.0696
0.097656 | -34.1419 | 8.1475
0.12207 | -33.9859 | 5.7839
0.14648 | -34.0107 | 3.9538
0.1709 | -34.0038 | 3.4654
0.19531 | -34.0163 | 2.478
0.21973 | -34.0084 | 2.4596
0.24414 | -34.0095 | 1.5867
0.26855 | -33.9734 | 1.673
0.29297 | -34.0042 | 1.3868
0.31738 | -33.9481 | 0.8043
0.3418 | -33.9352 | 0.60668
0.36621 | -33.959 | 0.40856
0.39062 | -33.9693 | -0.24541
0.41504 | -34.0098 | 0.5251
0.43945 | -33.9861 | -0.0023031
0.46387 | -33.9984 | -0.21188
0.48828 | -33.9392 | 0.20691
0.5127 | -34.0067 | -0.037568
0.53711 | -33.9984 | -0.27602
0.56152 | -33.9872 | -0.4436
0.58594 | -33.9196 | -0.70534
0.61035 | -34.0242 | -0.86976
0.63477 | -33.9619 | -0.80747
0.65918 | -34.0111 | -1.1224
0.68359 | -33.9828 | -0.75563
0.70801 | -34.0185 | -1.2597
0.73242 | -34.041 | -0.934
0.75684 | -34.0221 | -1.4148
0.78125 | -33.9606 | -1.0774
0.80566 | -34.0369 | -1.7376
0.83008 | -34.0597 | -1.2398
0.85449 | -33.9523 | -1.4332
0.87891 | -34.0358 | -1.7161
0.90332 | -34.0102 | -1.3822
0.92773 | -33.97 | -1.4622
0.95215 | -34.0589 | -1.7281
0.97656 | -34.0604 | -1.7712
1.001 | -34.0196 | -1.7451
1.0254 | -34.0377 | -1.875
1.0498 | -33.9919 | -1.8838
1.0742 | -33.9805 | -2.5758
1.0986 | -34.0007 | -2.3734
1.123 | -33.9638 | -2.5892
1.1475 | -34.0385 | -2.3058
1.1719 | -33.9993 | -2.6425
1.1963 | -33.9913 | -2.5899
1.2207 | -34.0095 | -3.0269
1.2451 | -34.009 | -2.7706
1.2695 | -33.9629 | -3.1052
1.2939 | -34.0308 | -2.6974
1.3184 | -34.016 | -3.0067
1.3428 | -34.0702 | -2.9701
1.3672 | -34.0486 | -3.3625
1.3916 | -34.0656 | -3.2746
1.416 | -34.005 | -3.1295
1.4404 | -34.0507 | -3.1234
1.4648 | -34.0068 | -3.0931
1.4893 | -34.0271 | -3.4736
1.5137 | -34.0578 | -3.1662
1.5381 | -34.0643 | -3.5855
1.5625 | -34.053 | -3.3625
1.5869 | -34.0171 | -3.6576
1.6113 | -34.0604 | -3.8823
1.6357 | -34.0846 | -4.2571
1.6602 | -34.0184 | -3.7766
1.6846 | -34.0276 | -4.1851
1.709 | -34.0497 | -3.8425
1.7334 | -34.036 | -4.0406
1.7578 | -34.0062 | -4.3043
1.7822 | -34.0402 | -4.1654
1.8066 | -34.049 | -4.2441
1.8311 | -34.0182 | -4.0344
1.8555 | -34.0485 | -4.189
1.8799 | -34.0173 | -4.6735
1.9043 | -34.0456 | -4.6995
1.9287 | -34.0289 | -4.5752
1.9531 | -34.0084 | -4.7186
1.9775 | -34.0235 | -4.5609
2.002 | -34.0667 | -4.9511
2.0264 | -34.019 | -4.6581
2.0508 | -34.0352 | -4.7319
2.0752 | -34.1149 | -4.6804
2.0996 | -34.0521 | -4.9294
2.124 | -34.0322 | -4.9551
2.1484 | -34.0499 | -5.4923
2.1729 | -34.0528 | -5.2579
2.1973 | -34.0353 | -5.3045
2.2217 | -34.0953 | -5.1512
2.2461 | -34.0447 | -5.4487
2.2705 | -34.0777 | -5.3722
2.2949 | -34.0145 | -5.2535
2.3193 | -34.0871 | -5.7175
2.3438 | -34.037 | -5.7366
2.3682 | -34.0432 | -5.6657
2.3926 | -34.0415 | -5.6103
2.417 | -34.0552 | -5.5891
2.4414 | -34.0295 | -5.8105
2.4658 | -34.1105 | -6.045
2.4902 | -34.0274 | -5.7768
2.5146 | -34.0496 | -6.0017
2.5391 | -34.0846 | -6.1062
2.5635 | -34.0816 | -6.1215
2.5879 | -34.0855 | -5.9412
2.6123 | -34.0246 | -6.422
2.6367 | -34.099 | -6.3399
2.6611 | -34.0508 | -6.5057
2.6855 | -34.0481 | -6.7051
2.71 | -34.0503 | -6.6579
2.7344 | -34.0738 | -6.6644
2.7588 | -34.0523 | -6.9561
2.7832 | -34.0818 | -6.7729
2.8076 | -34.1033 | -6.6106
2.832 | -34.0425 | -6.8002
2.8564 | -34.087 | -6.9699
2.8809 | -34.1075 | -7.1806
2.9053 | -34.0509 | -7.0085
2.9297 | -34.062 | -7.0886
2.9541 | -34.0903 | -7.2806
2.9785 | -34.0393 | -7.0872
3.0029 | -34.0338 | -7.3443
3.0273 | -34.0846 | -7.3975
3.0518 | -34.0931 | -7.3014
3.0762 | -34.0822 | -7.2678
3.1006 | -34.0916 | -7.2945
3.125 | -34.0849 | -7.4936
3.1494 | -34.0857 | -7.6246
3.1738 | -34.1181 | -7.7305
3.1982 | -34.1029 | -7.7346
3.2227 | -34.0707 | -7.7773
3.2471 | -34.1368 | -7.8502
3.2715 | -34.0766 | -7.6988
3.2959 | -34.1086 | -7.88
3.3203 | -34.0738 | -7.5049
3.3447 | -34.1479 | -8.13
3.3691 | -34.1218 | -7.8596
3.3936 | -34.0887 | -8.2378
3.418 | -34.1188 | -8.2502
3.4424 | -34.0786 | -8.4802
3.4668 | -34.0767 | -8.3349
3.4912 | -34.1025 | -8.6598
3.5156 | -34.1275 | -8.4022
3.54 | -34.0781 | -8.5406
3.5645 | -34.145 | -8.37
3.5889 | -34.0703 | -8.541
3.6133 | -34.1195 | -8.5154
3.6377 | -34.0739 | -8.7169
3.6621 | -34.1162 | -8.9302
3.6865 | -34.1058 | -8.6121
3.7109 | -34.0959 | -8.6122
3.7354 | -34.0809 | -8.9924
3.7598 | -34.1717 | -9.1199
3.7842 | -34.1348 | -8.7587
3.8086 | -34.0845 | -9.294
3.833 | -34.11 | -9.1402
3.8574 | -34.0751 | -9.1354
3.8818 | -34.1016 | -9.1055
3.9062 | -34.0703 | -9.5937
3.9307 | -34.1351 | -9.3404
3.9551 | -34.1314 | -9.334
3.9795 | -34.1833 | -9.488
4.0039 | -34.105 | -9.7295
4.0283 | -34.183 | -9.8975
4.0527 | -34.0826 | -9.5649
4.0771 | -34.1058 | -9.5157
4.1016 | -34.1388 | -10.1862
4.126 | -34.1815 | -10.0028
4.1504 | -34.1344 | -10.0501
4.1748 | -34.1082 | -10.0518
4.1992 | -34.1197 | -9.9618
4.2236 | -34.1339 | -10.2698
4.248 | -34.115 | -10.1505
4.2725 | -34.1293 | -10.1443
4.2969 | -34.0864 | -9.9497
4.3213 | -34.1545 | -10.3296
4.3457 | -34.1406 | -10.0556
4.3701 | -34.1004 | -10.6409
4.3945 | -34.1825 | -10.2158
4.4189 | -34.1055 | -10.3521
4.4434 | -34.205 | -10.494
4.4678 | -34.1382 | -10.7046
4.4922 | -34.1744 | -10.9188
4.5166 | -34.1163 | -10.6709
4.541 | -34.1215 | -10.9677
4.5654 | -34.1052 | -11.0232
4.5898 | -34.1514 | -11.1347
4.6143 | -34.1364 | -11.0672
4.6387 | -34.1462 | -11.0734
4.6631 | -34.1078 | -10.8941
4.6875 | -34.1422 | -11.4422
4.7119 | -34.1091 | -11.5561
4.7363 | -34.1317 | -11.6995
4.7607 | -34.1815 | -11.4826
4.7852 | -34.1544 | -11.6988
4.8096 | -34.159 | -11.5846
4.834 | -34.1431 | -11.8122
4.8584 | -34.1887 | -11.5944
4.8828 | -34.1809 | -11.7057
4.9072 | -34.1609 | -11.5434
4.9316 | -34.1416 | -12.0588
4.9561 | -34.1332 | -12.1129
4.9805 | -34.1537 | -12.1563
5.0049 | -34.1618 | -12.5018
5.0293 | -34.1125 | -12.1176
5.0537 | -34.1462 | -12.1996
5.0781 | -34.1673 | -11.9247
5.1025 | -34.151 | -12.2586
5.127 | -34.1114 | -12.2288
5.1514 | -34.1993 | -12.6095
5.1758 | -34.1751 | -12.1925
5.2002 | -34.1644 | -12.5342
5.2246 | -34.1983 | -12.509
5.249 | -34.099 | -12.9137
5.2734 | -34.206 | -12.2951
5.2979 | -34.1921 | -12.6479
5.3223 | -34.1827 | -12.8426
5.3467 | -34.121 | -13.0514
5.3711 | -34.2077 | -12.9772
5.3955 | -34.1834 | -13.1261
5.4199 | -34.1904 | -12.9842
5.4443 | -34.1598 | -13.3123
5.4688 | -34.2046 | -13.4162
5.4932 | -34.1062 | -13.0304
5.5176 | -34.1841 | -13.3055
5.542 | -34.145 | -13.6769
5.5664 | -34.1919 | -13.6216
5.5908 | -34.1938 | -13.4674
5.6152 | -34.1552 | -13.6038
5.6396 | -34.1854 | -13.8366
5.6641 | -34.1583 | -13.8534
5.6885 | -34.1414 | -13.3884
5.7129 | -34.1484 | -13.6226
5.7373 | -34.1495 | -13.3885
5.7617 | -34.1612 | -13.7581
5.7861 | -34.2092 | -13.8081
5.8105 | -34.2067 | -13.778
5.835 | -34.2275 | -13.5362
5.8594 | -34.2134 | -13.6502
5.8838 | -34.1766 | -14.3
5.9082 | -34.1334 | -14.3711
5.9326 | -34.2013 | -14.2809
5.957 | -34.1882 | -14.1376
5.9814 | -34.212 | -14.8482
6.0059 | -34.1976 | -14.4924
6.0303 | -34.1833 | -14.5341
6.0547 | -34.1825 | -14.3776
6.0791 | -34.2139 | -14.6745
6.1035 | -34.1857 | -14.204
6.1279 | -34.1991 | -14.4895
6.1523 | -34.1684 | -14.7566
6.1768 | -34.1828 | -15.0983
6.2012 | -34.2049 | -14.7859
6.2256 | -34.166 | -14.7862
6.25 | -34.2388 | -15.127
6.2744 | -34.1786 | -15.2074
6.2988 | -34.2313 | -14.9462
6.3232 | -34.2219 | -14.9271
6.3477 | -34.1962 | -15.1098
6.3721 | -34.1933 | -14.9555
6.3965 | -34.2401 | -15.1648
6.4209 | -34.1922 | -15.0662
6.4453 | -34.2337 | -15.6143
6.4697 | -34.2064 | -15.0501
6.4941 | -34.236 | -15.878
6.5186 | -34.2715 | -15.3268
6.543 | -34.2195 | -15.5602
6.5674 | -34.1959 | -15.6355
6.5918 | -34.1997 | -16.1985
6.6162 | -34.2295 | -15.7798
6.6406 | -34.2253 | -16.0304
6.665 | -34.2651 | -16.2434
6.6895 | -34.2696 | -15.7139
6.7139 | -34.268 | -16.3045
6.7383 | -34.2952 | -16.1688
6.7627 | -34.2309 | -15.9158
6.7871 | -34.1852 | -16.0503
6.8115 | -34.2592 | -15.8283
6.8359 | -34.2584 | -16.1193
6.8604 | -34.1868 | -16.1962
6.8848 | -34.2114 | -16.5591
6.9092 | -34.2056 | -16.4562
6.9336 | -34.2535 | -16.5911
6.958 | -34.2557 | -16.7842
6.9824 | -34.2046 | -16.4944
7.0068 | -34.2364 | -16.887
7.0312 | -34.3055 | -16.6235
7.0557 | -34.2756 | -16.9519
7.0801 | -34.2662 | -16.5049
7.1045 | -34.2591 | -16.9731
7.1289 | -34.2498 | -16.7468
7.1533 | -34.2024 | -17.4991
7.1777 | -34.2276 | -17.0219
7.2021 | -34.2233 | -17.2453
7.2266 | -34.2849 | -17.2985
7.251 | -34.2671 | -17.5007
7.2754 | -34.3108 | -17.4274
7.2998 | -34.2154 | -17.2928
7.3242 | -34.3076 | -17.279
7.3486 | -34.2971 | -17.6044
7.373 | -34.2624 | -17.5353
7.3975 | -34.2516 | -17.517
7.4219 | -34.3057 | -17.8279
7.4463 | -34.2902 | -17.7764
7.4707 | -34.2169 | -18.0797
7.4951 | -34.2665 | -17.6865
7.5195 | -34.2452 | -18.0059
7.5439 | -34.3033 | -17.826
7.5684 | -34.2965 | -17.9328
7.5928 | -34.3392 | -18.0534
7.6172 | -34.2581 | -17.9637
7.6416 | -34.3423 | -18.0952
7.666 | -34.2847 | -17.9683
7.6904 | -34.3012 | -18.3096
7.7148 | -34.3162 | -17.9581
7.7393 | -34.3221 | -18.3358
7.7637 | -34.2822 | -18.4171
7.7881 | -34.3256 | -18.8165
7.8125 | -34.2872 | -18.4505
7.8369 | -34.2691 | -18.7196
7.8613 | -34.2901 | -18.5848
7.8857 | -34.2798 | -18.5137
7.9102 | -34.3082 | -18.7766
7.9346 | -34.3058 | -19.0557
7.959 | -34.3242 | -18.7586
7.9834 | -34.2958 | -19.1132
8.0078 | -34.3541 | -18.9544
8.0322 | -34.2687 | -19.2715
8.0566 | -34.3286 | -19.0041
8.0811 | -34.3215 | -18.792
8.1055 | -34.3772 | -19.2233
8.1299 | -34.2715 | -19.1501
8.1543 | -34.2947 | -19.3284
8.1787 | -34.3128 | -19.1463
8.2031 | -34.3587 | -19.352
8.2275 | -34.3568 | -19.2761
8.252 | -34.3942 | -19.603
8.2764 | -34.3394 | -19.5939
8.3008 | -34.365 | -19.7469
8.3252 | -34.3827 | -19.562
8.3496 | -34.3132 | -19.7959
8.374 | -34.337 | -19.6749
8.3984 | -34.2973 | -20.12
8.4229 | -34.3687 | -20.031
8.4473 | -34.3635 | -20.5768
8.4717 | -34.4161 | -19.9974
8.4961 | -34.3282 | -20.2663
8.5205 | -34.322 | -20.0897
8.5449 | -34.3607 | -19.861
8.5693 | -34.358 | -20.1197
8.5938 | -34.3818 | -20.423
8.6182 | -34.3582 | -20.5248
8.6426 | -34.3109 | -20.5924
8.667 | -34.3718 | -20.7553
8.6914 | -34.3153 | -20.5985
8.7158 | -34.3897 | -20.7045
8.7402 | -34.3373 | -20.5076
8.7646 | -34.344 | -20.7871
8.7891 | -34.3212 | -20.7477
8.8135 | -34.3463 | -21.1821
8.8379 | -34.3825 | -20.6297
8.8623 | -34.3798 | -21.2762
8.8867 | -34.4089 | -20.8445
8.9111 | -34.3433 | -21.091
8.9355 | -34.4155 | -21.0973
8.96 | -34.321 | -21.2638
8.9844 | -34.35 | -21.2516
9.0088 | -34.3389 | -21.2821
9.0332 | -34.3991 | -21.5873
9.0576 | -34.3576 | -21.5123
9.082 | -34.383 | -21.2339
9.1064 | -34.4089 | -21.2638
9.1309 | -34.4459 | -21.6355
9.1553 | -34.3598 | -21.5295
9.1797 | -34.3646 | -21.5173
9.2041 | -34.4192 | -21.5698
9.2285 | -34.3923 | -22.3354
9.2529 | -34.3903 | -21.7113
9.2773 | -34.4062 | -21.8927
9.3018 | -34.3699 | -21.7182
9.3262 | -34.4657 | -22.0376
9.3506 | -34.4018 | -21.9674
9.375 | -34.4138 | -21.9424
9.3994 | -34.4166 | -22.0338
9.4238 | -34.4413 | -22.2227
9.4482 | -34.4426 | -22.2497
9.4727 | -34.4138 | -22.2305
9.4971 | -34.4367 | -22.3478
9.5215 | -34.4206 | -22.4817
9.5459 | -34.4282 | -22.6462
9.5703 | -34.3454 | -22.5147
9.5947 | -34.4256 | -22.6334
9.6191 | -34.4102 | -22.633
9.6436 | -34.4132 | -22.3906
9.668 | -34.4331 | -22.6775
9.6924 | -34.4051 | -22.869
9.7168 | -34.386 | -22.7199
9.7412 | -34.3704 | -23.0612
9.7656 | -34.3958 | -22.8636
9.79 | -34.4301 | -22.8694
9.8145 | -34.4624 | -23.3219
9.8389 | -34.4355 | -23.2618
9.8633 | -34.4153 | -22.8955
9.8877 | -34.3981 | -23.3191
9.9121 | -34.4577 | -23.0039
9.9365 | -34.4411 | -23.1392
9.9609 | -34.4241 | -23.0735
9.9854 | -34.4428 | -23.3534
10.0098 | -34.4533 | -23.5536

## ACOUS_TF
Frequency (Hz) | Amplitude (dB V/Pa) | Phase (degrees)
0.073242 | -34.1781 | 9.8444
0.097656 | -34.1924 | 5.1299
0.12207 | -34.009 | 3.6811
0.14648 | -34.0767 | 1.6146
0.1709 | -34.0048 | 1.1891
0.19531 | -34.0361 | 0.061545
0.21973 | -34.045 | 0.28184
0.24414 | -34.0468 | -0.93379
0.26855 | -34.0599 | -0.32425
0.29297 | -34.014 | -1.6068
0.31738 | -34.0582 | -1.1079
0.3418 | -33.9903 | -1.8238
0.36621 | -34.0804 | -1.5235
0.39062 | -33.9787 | -2.2181
0.41504 | -34.0987 | -1.7956
0.43945 | -33.9935 | -2.2749
0.46387 | -34.0943 | -2.4315
0.48828 | -33.9629 | -2.4354
0.5127 | -34.125 | -2.7793
0.53711 | -34.0045 | -2.4756
0.56152 | -34.0981 | -3.0968
0.58594 | -34.0124 | -2.6187
0.61035 | -34.0859 | -3.4066
0.63477 | -34.0373 | -2.7388
0.65918 | -34.0756 | -3.6401
0.68359 | -34.0346 | -2.8868
0.70801 | -34.0398 | -3.7432
0.73242 | -34.0778 | -3.073
0.75684 | -34.0448 | -3.9514
0.78125 | -34.1003 | -3.2803
0.80566 | -34.0245 | -4.0714
0.83008 | -34.1231 | -3.6301
0.85449 | -33.9955 | -4.0876
0.87891 | -34.0853 | -3.8244
0.90332 | -33.9859 | -3.9012
0.92773 | -34.1305 | -3.914
0.95215 | -34.0289 | -3.9186
0.97656 | -34.1326 | -4.4058
1.001 | -34.0271 | -4.0442
1.0254 | -34.1294 | -4.6314
1.0498 | -34.0684 | -4.0957
1.0742 | -34.1236 | -4.9906
1.0986 | -34.0479 | -4.2888
1.123 | -34.0719 | -4.9959
1.1475 | -34.0747 | -4.2332
1.1719 | -34.0669 | -5.1606
1.1963 | -34.101 | -4.2818
1.2207 | -34.0559 | -5.2848
1.2451 | -34.1329 | -4.5251
1.2695 | -34.033 | -5.1929
1.2939 | -34.134 | -4.7011
1.3184 | -34.0039 | -5.1165
1.3428 | -34.1655 | -4.9024
1.3672 | -34.0395 | -5.1079
1.3916 | -34.1447 | -5.275
1.416 | -34.0057 | -5.2448
1.4404 | -34.1334 | -5.4361
1.4648 | -34.0203 | -4.937
1.4893 | -34.1367 | -5.4837
1.5137 | -34.0676 | -5.0107
1.5381 | -34.1397 | -5.946
1.5625 | -34.0994 | -5.1268
1.5869 | -34.1216 | -5.9868
1.6113 | -34.1214 | -5.3535
1.6357 | -34.0689 | -6.3038
1.6602 | -34.1119 | -5.4978
1.6846 | -34.0526 | -6.1239
1.709 | -34.1375 | -5.6028
1.7334 | -34.026 | -6.1616
1.7578 | -34.1679 | -5.8985
1.7822 | -34.0246 | -6.066
1.8066 | -34.161 | -5.9748
1.8311 | -34.0325 | -6.0078
1.8555 | -34.1873 | -6.229
1.8799 | -34.0562 | -6.0164
1.9043 | -34.16 | -6.6361
1.9287 | -34.038 | -6.085
1.9531 | -34.1354 | -6.7426
1.9775 | -34.0733 | -5.8641
2.002 | -34.1238 | -6.8631
2.0264 | -34.0923 | -6.0472
2.0508 | -34.1073 | -6.7505
2.0752 | -34.1285 | -6.1589
2.0996 | -34.1094 | -6.9849
2.124 | -34.1613 | -6.2928
2.1484 | -34.0904 | -7.2463
2.1729 | -34.1632 | -6.6479
2.1973 | -34.0713 | -6.9761
2.2217 | -34.1694 | -6.7496
2.2461 | -34.0711 | -7.0101
2.2705 | -34.2067 | -6.9898
2.2949 | -34.0752 | -7.0862
2.3193 | -34.1816 | -7.2214
2.3438 | -34.085 | -6.9492
2.3682 | -34.1494 | -7.4178
2.3926 | -34.0749 | -6.7842
2.417 | -34.1429 | -7.5518
2.4414 | -34.1037 | -6.7613
2.4658 | -34.163 | -7.687
2.4902 | -34.1529 | -6.9278
2.5146 | -34.1292 | -7.962
2.5391 | -34.1441 | -7.1853
2.5635 | -34.1169 | -7.8974
2.5879 | -34.1739 | -7.2989
2.6123 | -34.0883 | -7.7903
2.6367 | -34.1836 | -7.4356
2.6611 | -34.0841 | -7.9036
2.6855 | -34.2075 | -7.6094
2.71 | -34.0672 | -7.8684
2.7344 | -34.2172 | -7.8943
2.7588 | -34.0953 | -7.5839
2.7832 | -34.1871 | -8.207
2.8076 | -34.0806 | -7.6372
2.832 | -34.197 | -8.2532
2.8564 | -34.1375 | -7.8058
2.8809 | -34.1663 | -8.6399
2.9053 | -34.1201 | -7.7751
2.9297 | -34.1562 | -8.6515
2.9541 | -34.1698 | -7.9267
2.9785 | -34.126 | -8.7144
3.0029 | -34.1773 | -7.9475
3.0273 | -34.1217 | -8.7001
3.0518 | -34.213 | -8.0419
3.0762 | -34.1092 | -8.6515
3.1006 | -34.2114 | -8.3197
3.125 | -34.1033 | -8.6151
3.1494 | -34.2179 | -8.5486
3.1738 | -34.1073 | -8.701
3.1982 | -34.2223 | -8.8519
3.2227 | -34.0827 | -8.577
3.2471 | -34.1951 | -9.1032
3.2715 | -34.1257 | -8.5751
3.2959 | -34.2043 | -9.0101
3.3203 | -34.1295 | -8.4614
3.3447 | -34.1694 | -9.3586
3.3691 | -34.1718 | -8.5958
3.3936 | -34.1816 | -9.3061
3.418 | -34.1869 | -8.7142
3.4424 | -34.165 | -9.5989
3.4668 | -34.2313 | -8.8328
3.4912 | -34.1406 | -9.6553
3.5156 | -34.2573 | -9.0924
3.54 | -34.1108 | -9.4997
3.5645 | -34.2334 | -9.3057
3.5889 | -34.1333 | -9.4626
3.6133 | -34.2591 | -9.4232
3.6377 | -34.1205 | -9.5308
3.6621 | -34.2305 | -9.6391
3.6865 | -34.1234 | -9.2803
3.7109 | -34.2463 | -9.7367
3.7354 | -34.1671 | -9.4396
3.7598 | -34.2235 | -10.0163
3.7842 | -34.1742 | -9.592
3.8086 | -34.195 | -10.1625
3.833 | -34.1763 | -9.4336
3.8574 | -34.183 | -10.1454
3.8818 | -34.2263 | -9.5539
3.9062 | -34.1633 | -10.3709
3.9307 | -34.2096 | -9.68
3.9551 | -34.1678 | -10.3264
3.9795 | -34.2856 | -10.0412
4.0039 | -34.1586 | -10.3298
4.0283 | -34.2734 | -10.1873
4.0527 | -34.128 | -10.3487
4.0771 | -34.2737 | -10.3819
4.1016 | -34.1699 | -10.1855
4.126 | -34.2797 | -10.5864
4.1504 | -34.1601 | -10.2884
4.1748 | -34.2564 | -10.9319
4.1992 | -34.1911 | -10.3234
4.2236 | -34.2223 | -11.0644
4.248 | -34.1855 | -10.2696
4.2725 | -34.1904 | -10.947
4.2969 | -34.1991 | -10.1649
4.3213 | -34.2185 | -11.1108
4.3457 | -34.2239 | -10.4494
4.3701 | -34.1734 | -11.1924
4.3945 | -34.2473 | -10.4635
4.4189 | -34.1691 | -11.0659
4.4434 | -34.2908 | -10.6776
4.4678 | -34.1901 | -11.1691
4.4922 | -34.2816 | -11.0854
4.5166 | -34.144 | -11.0993
4.541 | -34.2837 | -11.2408
4.5654 | -34.1733 | -10.9795
4.5898 | -34.2902 | -11.5027
4.6143 | -34.1756 | -10.8384
4.6387 | -34.2576 | -11.6362
4.6631 | -34.2109 | -10.9698
4.6875 | -34.2451 | -11.84
4.7119 | -34.22 | -11.0112
4.7363 | -34.2276 | -11.5964
4.7607 | -34.2392 | -11.0862
4.7852 | -34.2367 | -11.9372
4.8096 | -34.2728 | -11.2388
4.834 | -34.1912 | -12.0152
4.8584 | -34.3039 | -11.2887
4.8828 | -34.1863 | -11.7407
4.9072 | -34.3156 | -11.5394
4.9316 | -34.1895 | -11.9774
4.9561 | -34.3118 | -11.9535
4.9805 | -34.1835 | -11.7968
5.0049 | -34.2934 | -12.1652
5.0293 | -34.1562 | -11.7581
5.0537 | -34.2752 | -12.3364
5.0781 | -34.1972 | -11.5888
5.1025 | -34.2935 | -12.4646
5.127 | -34.2176 | -11.7874
5.1514 | -34.2812 | -12.5957
5.1758 | -34.2607 | -11.858
5.2002 | -34.2672 | -12.6592
5.2246 | -34.2802 | -11.8484
5.249 | -34.1998 | -12.6596
5.2734 | -34.287 | -11.9119
5.2979 | -34.225 | -12.8066
5.3223 | -34.3433 | -12.2846
5.3467 | -34.1987 | -12.6293
5.3711 | -34.3399 | -12.4554
5.3955 | -34.1888 | -12.6401
5.4199 | -34.3546 | -12.7475
5.4443 | -34.2161 | -12.5289
5.4688 | -34.3075 | -13.0557
5.4932 | -34.1989 | -12.5157
5.5176 | -34.3068 | -13.0717
5.542 | -34.2364 | -12.5697
5.5664 | -34.3035 | -13.3079
5.5908 | -34.2455 | -12.6144
5.6152 | -34.2884 | -13.5077
5.6396 | -34.2598 | -12.8707
5.6641 | -34.2484 | -13.5515
5.6885 | -34.2993 | -12.7493
5.7129 | -34.2236 | -13.5388
5.7373 | -34.304 | -12.9302
5.7617 | -34.2355 | -13.4388
5.7861 | -34.3575 | -12.8568
5.8105 | -34.2128 | -13.5325
5.835 | -34.3889 | -13.3687
5.8594 | -34.2268 | -13.4155
5.8838 | -34.3655 | -13.5003
5.9082 | -34.2325 | -13.4702
5.9326 | -34.346 | -13.8715
5.957 | -34.2491 | -13.1882
5.9814 | -34.3456 | -13.9476
6.0059 | -34.2692 | -13.4485
6.0303 | -34.3162 | -14.1287
6.0547 | -34.2718 | -13.3982
6.0791 | -34.2829 | -13.9746
6.1035 | -34.2917 | -13.149
6.1279 | -34.2712 | -14.3175
6.1523 | -34.3259 | -13.6128
6.1768 | -34.2406 | -14.3881
6.2012 | -34.3368 | -13.8336
6.2256 | -34.2401 | -14.0568
6.25 | -34.3652 | -13.8033
6.2744 | -34.2582 | -14.0371
6.2988 | -34.3518 | -14.1711
6.3232 | -34.2458 | -13.9337
6.3477 | -34.3699 | -14.0934
6.3721 | -34.2407 | -13.8866
6.3965 | -34.3618 | -14.4337
6.4209 | -34.2658 | -14.0146
6.4453 | -34.3468 | -14.5286
6.4697 | -34.2925 | -13.9021
6.4941 | -34.3397 | -14.9048
6.5186 | -34.3346 | -14.0265
6.543 | -34.3321 | -15.0215
6.5674 | -34.3635 | -14.0807
6.5918 | -34.3061 | -15.1524
6.6162 | -34.3452 | -14.4076
6.6406 | -34.2468 | -14.9955
6.665 | -34.3996 | -14.5088
6.6895 | -34.2661 | -14.827
6.7139 | -34.4075 | -14.8388
6.7383 | -34.2624 | -14.9362
6.7627 | -34.372 | -15.0599
6.7871 | -34.2396 | -14.7972
6.8115 | -34.418 | -15.1549
6.8359 | -34.3072 | -15.0936
6.8604 | -34.3936 | -15.3922
6.8848 | -34.2759 | -14.7812
6.9092 | -34.3658 | -15.5266
6.9336 | -34.3515 | -14.9685
6.958 | -34.345 | -15.7083
6.9824 | -34.3537 | -14.9262
7.0068 | -34.3425 | -15.6905
7.0312 | -34.3643 | -14.8698
7.0557 | -34.3255 | -15.5578
7.0801 | -34.3781 | -14.9584
7.1045 | -34.3267 | -15.565
7.1289 | -34.4332 | -15.2986
7.1533 | -34.2955 | -15.646
7.1777 | -34.4185 | -15.434
7.2021 | -34.3137 | -15.6549
7.2266 | -34.4503 | -15.6605
7.251 | -34.3408 | -15.7232
7.2754 | -34.4206 | -15.8253
7.2998 | -34.295 | -15.2112
7.3242 | -34.378 | -16.0826
7.3486 | -34.3389 | -15.3667
7.373 | -34.3861 | -16.2199
7.3975 | -34.3661 | -15.447
7.4219 | -34.3926 | -16.4568
7.4463 | -34.3798 | -15.6084
7.4707 | -34.3414 | -16.3855
7.4951 | -34.3688 | -15.8354
7.5195 | -34.3435 | -16.3796
7.5439 | -34.4195 | -15.929
7.5684 | -34.325 | -16.2247
7.5928 | -34.4024 | -16.0719
7.6172 | -34.2911 | -16.2431
7.6416 | -34.4465 | -16.2593
7.666 | -34.3032 | -16.0865
7.6904 | -34.4725 | -16.5894
7.7148 | -34.3623 | -16.2627
7.7393 | -34.4417 | -16.7705
7.7637 | -34.3948 | -16.313
7.7881 | -34.443 | -17.1268
7.8125 | -34.3864 | -16.4138
7.8369 | -34.3893 | -17.1675
7.8613 | -34.3992 | -16.18
7.8857 | -34.3855 | -17.3086
7.9102 | -34.4445 | -16.3091
7.9346 | -34.3482 | -17.1718
7.959 | -34.4536 | -16.5103
7.9834 | -34.3587 | -17.1857
8.0078 | -34.4758 | -16.6358
8.0322 | -34.3712 | -17.1483
8.0566 | -34.4847 | -16.935
8.0811 | -34.3655 | -16.8972
8.1055 | -34.4779 | -16.9288
8.1299 | -34.3381 | -16.9395
8.1543 | -34.4576 | -17.4008
8.1787 | -34.3657 | -16.8687
8.2031 | -34.4248 | -17.3166
8.2275 | -34.3909 | -16.9137
8.252 | -34.4315 | -17.6721
8.2764 | -34.423 | -17.206
8.3008 | -34.4315 | -17.7868
8.3252 | -34.4513 | -17.2901
8.3496 | -34.4341 | -17.9768
8.374 | -34.4864 | -17.115
8.3984 | -34.3997 | -18.1652
8.4229 | -34.4523 | -17.3024
8.4473 | -34.4115 | -17.9895
8.4717 | -34.4828 | -17.481
8.4961 | -34.3765 | -18.0221
8.5205 | -34.5038 | -17.3601
8.5449 | -34.3588 | -17.9358
8.5693 | -34.5319 | -17.8738
8.5938 | -34.385 | -17.518
8.6182 | -34.5236 | -18.2004
8.6426 | -34.4013 | -17.5325
8.667 | -34.4746 | -18.303
8.6914 | -34.4055 | -17.2605
8.7158 | -34.4394 | -18.475
8.7402 | -34.4121 | -17.6253
8.7646 | -34.4518 | -18.407
8.7891 | -34.4922 | -17.765
8.8135 | -34.4397 | -18.6566
8.8379 | -34.52 | -17.6465
8.8623 | -34.4304 | -18.3166
8.8867 | -34.5369 | -18.177
8.9111 | -34.4117 | -18.3025
8.9355 | -34.5193 | -18.4156
8.96 | -34.4164 | -18.5866
8.9844 | -34.5323 | -18.5703
9.0088 | -34.4201 | -18.4775
9.0332 | -34.5336 | -18.5474
9.0576 | -34.3753 | -18.3
9.082 | -34.5471 | -18.7614
9.1064 | -34.4425 | -18.0885
9.1309 | -34.5128 | -18.9804
9.1553 | -34.4335 | -18.3061
9.1797 | -34.4773 | -19.1475
9.2041 | -34.4804 | -18.4377
9.2285 | -34.4766 | -19.1452
9.2529 | -34.4564 | -18.1951
9.2773 | -34.4195 | -19.405
9.3018 | -34.501 | -18.4198
9.3262 | -34.4624 | -19.1601
9.3506 | -34.5721 | -18.7776
9.375 | -34.4316 | -18.9547
9.3994 | -34.5474 | -18.6665
9.4238 | -34.4572 | -18.9271
9.4482 | -34.5872 | -18.7831
9.4727 | -34.4445 | -18.9997
9.4971 | -34.5465 | -19.3819
9.5215 | -34.4291 | -18.91
9.5459 | -34.5299 | -19.645
9.5703 | -34.4745 | -18.7539
9.5947 | -34.5091 | -19.7879
9.6191 | -34.4766 | -19.0528
9.6436 | -34.5409 | -19.7566
9.668 | -34.5174 | -19.0081
9.6924 | -34.4512 | -19.7425
9.7168 | -34.5226 | -19.127
9.7412 | -34.4724 | -19.8704
9.7656 | -34.5599 | -19.1086
9.79 | -34.4329 | -19.8541
9.8145 | -34.5698 | -19.2349
9.8389 | -34.4416 | -20.2258
9.8633 | -34.5845 | -19.5958
9.8877 | -34.4272 | -19.7184
9.9121 | -34.5667 | -19.6801
9.9365 | -34.4557 | -19.8459
9.9609 | -34.5903 | -19.8316
9.9854 | -34.4805 | -19.6482
10.0098 | -34.5999 | -20.3757

## SEALED_NOISE
dBPa/√Hz | Frequency
0.008853 | -31.5002
0.009193 | -33.4839
0.009534 | -33.5603
0.009874 | -30.775
0.010215 | -28.8961
0.010555 | -29.4805
0.010896 | -32.6194
0.011236 | -36.8675
0.011577 | -38.0842
0.011917 | -39.098
0.012258 | -39.4587
0.012598 | -37.6766
0.012939 | -32.5639
0.013279 | -28.815
0.01362 | -28.0913
0.01396 | -30.1747
0.014301 | -33.9426
0.014641 | -36.0572
0.014981 | -34.2848
0.015322 | -32.4291
0.015662 | -32.9689
0.016003 | -35.6969
0.016343 | -38.5216
0.016684 | -38.2575
0.017024 | -36.06
0.017365 | -35.3534
0.017705 | -36.6577
0.018046 | -38.5591
0.018386 | -37.6088
0.018727 | -36.5016
0.019067 | -35.9626
0.019408 | -37.2242
0.019748 | -38.325
0.020089 | -39.5002
0.020429 | -41.0095
0.02077 | -42.9617
0.02111 | -44.5478
0.021451 | -42.616
0.021791 | -38.7627
0.022132 | -35.881
0.022472 | -34.5115
0.022813 | -34.518
0.023153 | -35.5513
0.023494 | -37.178
0.023834 | -38.9934
0.024175 | -40.8956
0.024515 | -43.2389
0.024856 | -45.1191
0.025196 | -43.7052
0.025537 | -41.107
0.025877 | -39.8557
0.026218 | -40.0216
0.026558 | -40.0146
0.026899 | -38.6032
0.027239 | -37.0867
0.02758 | -36.4109
0.02792 | -36.6064
0.028261 | -37.3312
0.028601 | -38.477
0.028941 | -40.3101
0.029282 | -41.8781
0.029622 | -43.648
0.029963 | -45.0288
0.030303 | -45.4399
0.030644 | -44.6145
0.030984 | -43.2799
0.031325 | -42.0952
0.031665 | -41.3227
0.032006 | -41.1052
0.032346 | -41.2149
0.032687 | -41.4159
0.033027 | -41.7428
0.033368 | -42.5577
0.033708 | -43.5833
0.034049 | -43.9081
0.034389 | -42.9172
0.03473 | -41.6204
0.03507 | -40.9267
0.035411 | -41.0414
0.035751 | -41.5189
0.036092 | -41.7765
0.036432 | -41.7619
0.036773 | -41.9144
0.037113 | -42.4308
0.037454 | -43.4161
0.037794 | -44.4418
0.038135 | -44.5743
0.038475 | -43.4858
0.038816 | -42.1737
0.039156 | -41.9873
0.039497 | -42.1022
0.039837 | -42.7944
0.040178 | -43.9942
0.040518 | -45.5227
0.040859 | -47.0086
0.041199 | -48.0389
0.04154 | -48.1581
0.04188 | -47.4239
0.042221 | -46.4825
0.042561 | -45.8636
0.042902 | -45.8081
0.043242 | -46.3125
0.043582 | -47.3135
0.043923 | -48.4809
0.044263 | -49.2585
0.044604 | -49.4005
0.044944 | -49.0407
0.045285 | -48.3731
0.045625 | -47.8463
0.045966 | -47.5986
0.046306 | -47.6949
0.046647 | -48.0834
0.046987 | -48.5033
0.047328 | -48.5125
0.047668 | -48.0957
0.048009 | -47.5823
0.048349 | -47.314
0.04869 | -47.4143
0.04903 | -47.8523
0.049371 | -48.1406
0.049711 | -48.3309
0.050052 | -48.365
0.050392 | -48.3194
0.050733 | -48.2143
0.051073 | -48.1222
0.051414 | -48.1005
0.051754 | -48.2151
0.052095 | -48.518
0.052435 | -48.9829
0.052776 | -49.5057
0.053116 | -49.922
0.053457 | -50.0051
0.053797 | -49.8063
0.054138 | -49.4832
0.054478 | -49.0845
0.054819 | -48.6371
0.055159 | -48.1813
0.0555 | -47.8532
0.05584 | -47.668
0.056181 | -47.6104
0.056521 | -47.5785
0.056862 | -47.518
0.057202 | -47.4785
0.057542 | -47.5384
0.057883 | -47.7015
0.058223 | -47.9346
0.058564 | -48.1724
0.058904 | -48.2102
0.059245 | -48.2883
0.059585 | -48.3423
0.059926 | -48.4613
0.060266 | -48.654
0.060607 | -48.9051
0.060947 | -49.1816
0.061288 | -49.4416
0.061628 | -49.6483
0.061969 | -49.8097
0.062309 | -50.0207
0.06265 | -50.4359
0.06299 | -50.9923
0.063331 | -51.4826
0.063671 | -51.8903
0.064012 | -52.062
0.064352 | -51.8384
0.064693 | -51.3429
0.065033 | -50.7451
0.065374 | -50.1785
0.065714 | -49.8465
0.066055 | -49.7526
0.066395 | -49.8999
0.066736 | -50.2863
0.067076 | -50.8517
0.067417 | -51.4706
0.067757 | -51.9918
0.068098 | -52.3552
0.068438 | -52.2957
0.068779 | -52.428
0.069119 | -52.4461
0.06946 | -52.4099
0.0698 | -52.4503
0.070141 | -52.6029
0.070481 | -52.7867
0.070822 | -53.059
0.071162 | -53.4017
0.071503 | -53.7955
0.071843 | -54.205
0.072183 | -54.4353
0.072524 | -54.6404
0.072864 | -54.6919
0.073205 | -54.5023
0.073545 | -54.1366
0.073886 | -53.6627
0.074226 | -53.2075
0.074567 | -52.874
0.074907 | -52.6948
0.075248 | -52.6929
0.075588 | -52.9088
0.075929 | -53.294
0.076269 | -53.8307
0.07661 | -54.3989
0.07695 | -54.9384
0.077291 | -55.3763
0.077631 | -55.6634
0.077972 | -55.9112
0.078312 | -55.8786
0.078653 | -55.8731
0.078993 | -55.9125
0.079334 | -55.9713
0.079674 | -56.0029
0.080015 | -56.1051
0.080355 | -56.2524
0.080696 | -56.393
0.081036 | -56.5538
0.081377 | -56.6959
0.081717 | -56.7262
0.082058 | -56.6666
0.082398 | -56.4892
0.082739 | -56.2117
0.083079 | -55.8674
0.08342 | -55.4969
0.08376 | -55.1401
0.084101 | -54.7548
0.084441 | -54.4454
0.084782 | -54.2448
0.085122 | -54.1133
0.085463 | -54.0411
0.085803 | -54.0076
0.086143 | -54.0306
0.086484 | -54.0955
0.086824 | -54.2269
0.087165 | -54.4663
0.087505 | -54.8095
0.087846 | -55.2919
0.088186 | -55.9231
0.088527 | -56.5709
0.088867 | -57.2351
0.089208 | -57.8793
0.089548 | -58.4588
0.089889 | -58.846
0.090229 | -58.926
0.09057 | -58.6869
0.09091 | -58.2898
0.091251 | -57.7402
0.091591 | -57.0597
0.091932 | -56.4251
0.092272 | -55.8501
0.092613 | -55.3439
0.092953 | -54.8837
0.093294 | -54.5425
0.093634 | -54.3263
0.093975 | -54.1897
0.094315 | -54.1311
0.094656 | -54.1046
0.094996 | -54.0991
0.095337 | -54.1257
0.095677 | -54.1643
0.096018 | -54.2164
0.096358 | -54.2718
0.096699 | -54.3436
0.097039 | -54.449
0.09738 | -54.6223
0.09772 | -54.8466
0.098061 | -55.1205
0.098401 | -55.3736
0.098742 | -55.6233
0.099082 | -55.9163
0.099423 | -56.25
0.099763 | -56.6084
0.1001 | -56.9985
0.10044 | -57.4402
0.10078 | -57.8069
0.10113 | -58.0857
0.10146 | -58.3469
0.10181 | -58.5282
0.10215 | -58.652
0.10249 | -58.6622
0.10283 | -58.5181
0.10317 | -58.3007
0.10351 | -58.0638
0.10385 | -57.8399
0.10419 | -57.6139
0.10453 | -57.4171
0.10487 | -57.2659
0.10521 | -57.1843
0.10555 | -57.1936
0.10589 | -57.275
0.10623 | -57.4265
0.10657 | -57.6412
0.10691 | -57.8714
0.10725 | -58.1141
0.10759 | -58.3865
0.10794 | -58.6425
0.10827 | -58.9362
0.10862 | -59.201
0.10896 | -59.4615
0.1093 | -59.7612
0.10964 | -60.0103
0.10998 | -60.2464
0.11032 | -60.4359
0.11066 | -60.5781
0.111 | -60.6536
0.11134 | -60.6756
0.11168 | -60.628
0.11202 | -60.5578
0.11236 | -60.504
0.1127 | -60.4725
0.11304 | -60.4246
0.11338 | -60.4257
0.11372 | -60.4226
0.11406 | -60.4665
0.1144 | -60.5536
0.11474 | -60.6642
0.11509 | -60.7929
0.11542 | -60.9194
0.11577 | -61.0048
0.11611 | -61.1049
0.11645 | -61.151
0.11679 | -61.1746
0.11713 | -61.2198
0.11747 | -61.2521
0.11781 | -61.2576
0.11815 | -61.2549
0.11849 | -61.2896
0.11883 | -61.3207
0.11917 | -61.3438
0.11951 | -61.3532
0.11985 | -61.3705
0.12019 | -61.363
0.12053 | -61.3852
0.12087 | -61.3791
0.12121 | -61.3041
0.12155 | -61.2112
0.1219 | -61.0876
0.12223 | -60.9069
0.12258 | -60.7285
0.12292 | -60.5259
0.12326 | -60.3402
0.1236 | -60.1979
0.12394 | -60.0897
0.12428 | -60.0391
0.12462 | -60.0309
0.12496 | -60.043
0.1253 | -60.1327
0.12564 | -60.2468
0.12598 | -60.3633
0.12632 | -60.4843
0.12666 | -60.5951
0.127 | -60.6969
0.12734 | -60.774
0.12768 | -60.8181
0.12802 | -60.8811
0.12836 | -60.9618
0.12871 | -61.0529
0.12904 | -61.1439
0.12939 | -61.2443
0.12973 | -61.3714
0.13007 | -61.4826
0.13041 | -61.6038
0.13075 | -61.7422
0.13109 | -61.8456
0.13143 | -61.956
0.13177 | -62.0324
0.13211 | -62.0895
0.13245 | -62.1414
0.13279 | -62.1804
0.13313 | -62.2141
0.13347 | -62.2698
0.13381 | -62.3262
0.13415 | -62.4213
0.13449 | -62.543
0.13483 | -62.6528
0.13517 | -62.7924
0.13551 | -62.9418
0.13586 | -63.1213
0.1362 | -63.3168
0.13654 | -63.4493
0.13688 | -63.5499
0.13722 | -63.5666
0.13756 | -63.6266
0.1379 | -63.6707
0.13824 | -63.642
0.13858 | -63.4992
0.13892 | -63.3246
0.13926 | -63.1558
0.1396 | -62.9282
0.13994 | -62.6904
0.14028 | -62.4556
0.14062 | -62.2128
0.14096 | -61.9883
0.1413 | -61.8024
0.14164 | -61.6356
0.14198 | -61.5159
0.14232 | -61.4371
0.14266 | -61.3952
0.143 | -61.3906
0.14335 | -61.442
0.14369 | -61.5243
0.14403 | -61.6511
0.14437 | -61.8237
0.14471 | -62.0403
0.14505 | -62.2977
0.14539 | -62.587
0.14573 | -62.9079
0.14607 | -63.2212
0.14641 | -63.551
0.14675 | -63.9404
0.14709 | -64.1806
0.14743 | -64.5316
0.14777 | -64.7957
0.14811 | -65.0192
0.14845 | -65.1904
0.14879 | -65.3196
0.14913 | -65.4728
0.14947 | -65.6006
0.14982 | -65.6536
0.15016 | -65.6292
0.1505 | -65.5569
0.15084 | -65.4686
0.15118 | -65.4012
0.15152 | -65.3216
0.15186 | -65.2044
0.1522 | -65.0862
0.15254 | -64.9644
0.15288 | -64.8549
0.15322 | -64.7533
0.15356 | -64.6449
0.1539 | -64.5128
0.15424 | -64.3484
0.15458 | -64.1877
0.15492 | -64.0437
0.15526 | -63.8991
0.1556 | -63.745
0.15594 | -63.5922
0.15628 | -63.4446
0.15662 | -63.3069
0.15696 | -63.2082
0.15731 | -63.1207
0.15765 | -63.0438
0.15799 | -62.9797
0.15833 | -62.9391
0.15867 | -62.9239
0.15901 | -62.9416
0.15935 | -63.0055
0.15969 | -63.1115
0.16003 | -63.2494
0.16037 | -63.4172
0.16071 | -63.6027
0.16105 | -63.8314
0.16139 | -64.0803
0.16173 | -64.3471
0.16207 | -64.6261
0.16241 | -64.8731
0.16275 | -65.0979
0.16309 | -65.3448
0.16343 | -65.5496
0.16378 | -65.6905
0.16412 | -65.7907
0.16446 | -65.8554
0.1648 | -65.8765
0.16514 | -65.8744
0.16548 | -65.8796
0.16582 | -65.8658
0.16616 | -65.8219
0.1665 | -65.7258
0.16684 | -65.657
0.16718 | -65.5681
0.16752 | -65.4974
0.16786 | -65.4416
0.1682 | -65.3863
0.16854 | -65.3546
0.16888 | -65.3496
0.16922 | -65.355
0.16956 | -65.3739
0.1699 | -65.4139
0.17024 | -65.4671
0.17058 | -65.5319
0.17092 | -65.609
0.17127 | -65.6883
0.17161 | -65.7667
0.17195 | -65.8522
0.17229 | -65.9565
0.17263 | -66.046
0.17297 | -66.137
0.17331 | -66.2174
0.17365 | -66.2877
0.17399 | -66.3617
0.17433 | -66.4427
0.17467 | -66.531
0.17501 | -66.6064
0.17535 | -66.6529
0.17569 | -66.7369
0.17603 | -66.8099
0.17637 | -66.8569
0.17671 | -66.9134
0.17705 | -66.9571
0.17739 | -67.0166
0.17774 | -67.0803
0.17808 | -67.1486
0.17842 | -67.2388
0.17876 | -67.3286
0.1791 | -67.4324
0.17944 | -67.5286
0.17978 | -67.6295
0.18012 | -67.7636
0.18046 | -67.9055
0.1808 | -68.0612
0.18114 | -68.2403
0.18148 | -68.427
0.18182 | -68.5805
0.18216 | -68.7337
0.1825 | -68.909
0.18284 | -69.0747
0.18318 | -69.2055
0.18352 | -69.332
0.18386 | -69.4462
0.1842 | -69.4985
0.18454 | -69.4825
0.18488 | -69.4415
0.18523 | -69.3656
0.18557 | -69.2722
0.18591 | -69.1249
0.18625 | -68.9833
0.18659 | -68.8186
0.18693 | -68.7075
0.18727 | -68.5923
0.18761 | -68.4828
0.18795 | -68.3905
0.18829 | -68.29
0.18863 | -68.1875
0.18897 | -68.1218
0.18931 | -68.0833
0.18965 | -68.0638
0.18999 | -68.063
0.19033 | -68.0918
0.19067 | -68.1377
0.19101 | -68.1936
0.19135 | -68.2681
0.1917 | -68.3617
0.19204 | -68.4656
0.19238 | -68.5743
0.19272 | -68.6924
0.19306 | -68.8321
0.1934 | -68.978
0.19374 | -69.108
0.19408 | -69.2443
0.19442 | -69.3883
0.19476 | -69.5358
0.1951 | -69.6474
0.19544 | -69.7937
0.19578 | -69.9363
0.19612 | -69.9908
0.19646 | -69.9377
0.1968 | -69.9635
0.19714 | -69.9781
0.19748 | -69.9673
0.19782 | -69.9411
0.19816 | -69.9061
0.1985 | -69.8094
0.19884 | -69.6701
0.19919 | -69.5305
0.19953 | -69.3536
0.19987 | -69.1758
0.20021 | -69.0065
0.20055 | -68.8035
0.20089 | -68.5908
0.20123 | -68.4153
0.20157 | -68.2427
0.20191 | -68.0783
0.20225 | -67.9333
0.20259 | -67.8124
0.20293 | -67.7132
0.20327 | -67.6113
0.20361 | -67.5143
0.20395 | -67.4463
0.20429 | -67.4016
0.20463 | -67.3713
0.20497 | -67.3633
0.20531 | -67.3717
0.20566 | -67.3947
0.206 | -67.5196
0.20634 | -67.5758
0.20668 | -67.6544
0.20702 | -67.7573
0.20736 | -67.866
0.2077 | -67.9722
0.20804 | -68.0967
0.20838 | -68.2376
0.20872 | -68.3873
0.20906 | -68.5457
0.2094 | -68.7135
0.20974 | -68.897
0.21008 | -69.0571
0.21042 | -69.21
0.21076 | -69.3811
0.2111 | -69.5339
0.21144 | -69.6419
0.21178 | -69.7242
0.21212 | -69.8018
0.21246 | -69.8659
0.2128 | -69.9134
0.21315 | -69.9631
0.21349 | -70.025
0.21383 | -70.0924
0.21417 | -70.1295
0.21451 | -70.1894
0.21485 | -70.2455
0.21519 | -70.2856
0.21553 | -70.3596
0.21587 | -70.4069
0.21621 | -70.4898
0.21655 | -70.5289
0.21689 | -70.5939
0.21723 | -70.6706
0.21757 | -70.7444
0.21791 | -70.8123
0.21825 | -70.8704
0.21859 | -70.9444
0.21893 | -71.0355
0.21927 | -71.0851
0.21962 | -71.1051
0.21996 | -71.129
0.2203 | -71.1427
0.22064 | -71.1626
0.22098 | -71.1851
0.22132 | -71.1819
0.22166 | -71.1743
0.222 | -71.1763
0.22234 | -71.1724
0.22268 | -71.1585
0.22302 | -71.1522
0.22336 | -71.1753
0.2237 | -71.1749
0.22404 | -71.1635
0.22438 | -71.1502
0.22472 | -71.1267
0.22506 | -71.0925
0.2254 | -71.0535
0.22574 | -71.0262
0.22608 | -71.0111
0.22642 | -70.9963
0.22676 | -70.9702
0.22711 | -70.9543
0.22745 | -70.9283
0.22779 | -70.9081
0.22813 | -70.8779
0.22847 | -70.8301
0.22881 | -70.7935
0.22915 | -70.7246
0.22949 | -70.6515
0.22983 | -70.5938
0.23017 | -70.5322
0.23051 | -70.463
0.23085 | -70.4011
0.23119 | -70.3431
0.23153 | -70.2935
0.23187 | -70.271
0.23221 | -70.2469
0.23255 | -70.208
0.23289 | -70.1827
0.23323 | -70.1511
0.23358 | -70.1326
0.23392 | -70.1244
0.23426 | -70.1315
0.2346 | -70.1531
0.23494 | -70.1754
0.23528 | -70.2101
0.23562 | -70.3132
0.23596 | -70.3847
0.2363 | -70.4736
0.23664 | -70.5486
0.23698 | -70.6382
0.23732 | -70.7481
0.23766 | -70.8578
0.238 | -70.9719
0.23834 | -71.1101
0.23868 | -71.2671
0.23902 | -71.4274
0.23936 | -71.5947
0.2397 | -71.7634
0.24004 | -71.9312
0.24038 | -72.1067
0.24072 | -72.273
0.24107 | -72.4399
0.24141 | -72.6094
0.24175 | -72.7753
0.24209 | -72.9305
0.24243 | -73.0418
0.24277 | -73.1075
0.24311 | -73.1592
0.24345 | -73.1947
0.24379 | -73.2315
0.24413 | -73.2463
0.24447 | -73.2354
0.24481 | -73.2165
0.24515 | -73.2104
0.24549 | -73.1796
0.24583 | -73.1728
0.24617 | -73.1284
0.24651 | -73.1354
0.24685 | -73.1428
0.24719 | -73.1174
0.24754 | -73.099
0.24788 | -73.0851
0.24822 | -73.056
0.24856 | -73.0339
0.2489 | -73.0049
0.24924 | -72.9755
0.24958 | -72.9648
0.24992 | -72.948
0.25026 | -72.9192
0.2506 | -72.895
0.25094 | -72.8527
0.25128 | -72.8093
0.25162 | -72.7571
0.25196 | -72.7062
0.2523 | -72.6438
0.25264 | -72.5672
0.25298 | -72.4973
0.25332 | -72.414
0.25366 | -72.3251
0.254 | -72.2366
0.25434 | -72.1405
0.25468 | -72.0215
0.25503 | -71.9016
0.25537 | -71.79
0.25571 | -71.6997
0.25605 | -71.6146
0.25639 | -71.5169
0.25673 | -71.4166
0.25707 | -71.3021
0.25741 | -71.2214
0.25775 | -71.1463
0.25809 | -71.0196
0.25843 | -70.9098
0.25877 | -70.8218
0.25911 | -70.7382
0.25945 | -70.6689
0.25979 | -70.6045
0.26013 | -70.543
0.26047 | -70.4927
0.26081 | -70.4467
0.26115 | -70.3983
0.26149 | -70.3473
0.26183 | -70.2931
0.26218 | -70.2436
0.26252 | -70.1961
0.26286 | -70.1552
0.2632 | -70.1158
0.26354 | -70.0795
0.26388 | -70.0448
0.26422 | -70.0187
0.26456 | -70.0019
0.2649 | -69.9806
0.26524 | -69.9952
0.26558 | -69.9787
0.26592 | -69.9627
0.26626 | -69.9571
0.2666 | -69.9702
0.26694 | -69.991
0.26728 | -70.0047
0.26762 | -70.0363
0.26796 | -70.0727
0.26831 | -70.0958
0.26865 | -70.1369
0.26899 | -70.1784
0.26933 | -70.2151
0.26967 | -70.275
0.27001 | -70.3214
0.27035 | -70.3706
0.27069 | -70.4203
0.27103 | -70.4699
0.27137 | -70.5183
0.27171 | -70.5424
0.27205 | -70.5497
0.27239 | -70.5661
0.27273 | -70.5855
0.27307 | -70.6051
0.27341 | -70.6182
0.27375 | -70.6038
0.27409 | -70.5732
0.27443 | -70.5468
0.27477 | -70.4864
0.27511 | -70.447
0.27546 | -70.4091
0.2758 | -70.3733
0.27614 | -70.3338
0.27648 | -70.3095
0.27682 | -70.2656
0.27716 | -70.2244
0.2775 | -70.2127
0.27784 | -70.1825
0.27818 | -70.1466
0.27852 | -70.1029
0.27886 | -70.0585
0.2792 | -70.0229
0.27954 | -69.9915
0.27988 | -69.9654
0.28022 | -69.95
0.28056 | -69.9392
0.2809 | -69.9381
0.28124 | -69.9491
0.28158 | -69.9521
0.28192 | -69.9495
0.28226 | -69.95
0.2826 | -69.9586
0.28295 | -69.972
0.28329 | -69.9903
0.28363 | -70.0083
0.28397 | -70.0283
0.28431 | -70.0525
0.28465 | -70.0868
0.28499 | -70.116
0.28533 | -70.1468
0.28567 | -70.1906
0.28601 | -70.2357
0.28635 | -70.277
0.28669 | -70.3161
0.28703 | -70.3578
0.28737 | -70.4081
0.28771 | -70.4501
0.28805 | -70.491
0.28839 | -70.5445
0.28873 | -70.6025
0.28907 | -70.6756
0.28941 | -70.7526
0.28975 | -70.8147
0.2901 | -70.8923
0.29044 | -70.9735
0.29078 | -71.055
0.29112 | -71.1419
0.29146 | -71.2271
0.2918 | -71.3289
0.29214 | -71.4381
0.29248 | -71.5346
0.29282 | -71.6247
0.29316 | -71.7192
0.2935 | -71.8177
0.29384 | -71.9163
0.29418 | -71.9971
0.29452 | -72.057
0.29486 | -72.1371
0.2952 | -72.2169
0.29554 | -72.2989
0.29588 | -72.371
0.29623 | -72.434
0.29657 | -72.4931
0.29691 | -72.5554
0.29725 | -72.6077
0.29759 | -72.652
0.29793 | -72.7046
0.29827 | -72.7505
0.29861 | -72.7741
0.29895 | -72.7916
0.29929 | -72.8306
0.29963 | -72.8608
0.29997 | -72.8818
0.30031 | -72.906
0.30065 | -72.9302
0.30099 | -72.9517
0.30133 | -72.978
0.30167 | -73.0013
0.30201 | -73.0157
0.30235 | -73.0504
0.30269 | -73.0965
0.30303 | -73.1275
0.30338 | -73.1507
0.30372 | -73.1788
0.30406 | -73.2121
0.3044 | -73.2585
0.30474 | -73.2813
0.30508 | -73.3156
0.30542 | -73.3477
0.30576 | -73.3815
0.3061 | -73.4159
0.30644 | -73.4441
0.30678 | -73.4713
0.30712 | -73.4997
0.30746 | -73.5255
0.3078 | -73.5434
0.30814 | -73.563
0.30848 | -73.5842
0.30882 | -73.5975
0.30916 | -73.6136
0.3095 | -73.6382
0.30984 | -73.662
0.31018 | -73.6548
0.31052 | -73.6483
0.31087 | -73.6316
0.31121 | -73.5896
0.31155 | -73.5574
0.31189 | -73.5272
0.31223 | -73.4869
0.31257 | -73.4376
0.31291 | -73.4042
0.31325 | -73.3634
0.31359 | -73.3173
0.31393 | -73.2524
0.31427 | -73.1885
0.31461 | -73.1291
0.31495 | -73.0686
0.31529 | -73.0092
0.31563 | -72.9634
0.31597 | -72.9256
0.31631 | -72.8804
0.31665 | -72.8247
0.31699 | -72.7806
0.31733 | -72.7413
0.31767 | -72.7035
0.31802 | -72.6797
0.31836 | -72.6608
0.3187 | -72.6451
0.31904 | -72.6374
0.31938 | -72.6266
0.31972 | -72.6224
0.32006 | -72.624
0.3204 | -72.6374
0.32074 | -72.6714
0.32108 | -72.6957
0.32142 | -72.7122
0.32176 | -72.7401
0.3221 | -72.7741
0.32244 | -72.8099
0.32278 | -72.8467
0.32312 | -72.8873
0.32346 | -72.931
0.3238 | -72.9715
0.32415 | -72.9932
0.32449 | -73.0247
0.32483 | -73.0459
0.32517 | -73.0706
0.32551 | -73.1005
0.32585 | -73.1101
0.32619 | -73.1048
0.32653 | -73.0906
0.32687 | -73.0657
0.32721 | -73.0462
0.32755 | -73.0287
0.32789 | -72.9973
0.32823 | -72.9571
0.32857 | -72.9109
0.32891 | -72.8667
0.32925 | -72.8229
0.32959 | -72.771
0.32993 | -72.7195
0.33027 | -72.6552
0.33061 | -72.5987
0.33095 | -72.5424
0.3313 | -72.4719
0.33164 | -72.4069
0.33198 | -72.3439
0.33232 | -72.2755
0.33266 | -72.2025
0.333 | -72.138
0.33334 | -72.0754
0.33368 | -72.0429
0.33402 | -71.9967
0.33436 | -71.9573
0.3347 | -71.9098
0.33504 | -71.8609
0.33538 | -71.8213
0.33572 | -71.7901
0.33606 | -71.7537
0.3364 | -71.7191
0.33674 | -71.691
0.33708 | -71.6694
0.33742 | -71.6528
0.33776 | -71.641
0.3381 | -71.6312
0.33844 | -71.6336
0.33879 | -71.6398
0.33913 | -71.6485
0.33947 | -71.6676
0.33981 | -71.6821
0.34015 | -71.7065
0.34049 | -71.7292
0.34083 | -71.7574
0.34117 | -71.7901
0.34151 | -71.8214
0.34185 | -71.8603
0.34219 | -71.9017
0.34253 | -71.9377
0.34287 | -71.9907
0.34321 | -72.047
0.34355 | -72.1059
0.34389 | -72.153
0.34423 | -72.1987
0.34457 | -72.2489
0.34491 | -72.3033
0.34525 | -72.3606
0.34559 | -72.4132
0.34594 | -72.4577
0.34628 | -72.5133
0.34662 | -72.5712
0.34696 | -72.627
0.3473 | -72.6853
0.34764 | -72.7415
0.34798 | -72.7977
0.34832 | -72.853
0.34866 | -72.9144
0.349 | -72.979
0.34934 | -73.0341
0.34968 | -73.0794
0.35002 | -73.1374
0.35036 | -73.2039
0.3507 | -73.2711
0.35104 | -73.3393
0.35138 | -73.398
0.35172 | -73.4456
0.35207 | -73.4891
0.35241 | -73.5326
0.35275 | -73.5642
0.35309 | -73.5914
0.35343 | -73.6093
0.35377 | -73.6335
0.35411 | -73.647
0.35445 | -73.6383
0.35479 | -73.6508
0.35513 | -73.6615
0.35547 | -73.6617
0.35581 | -73.6546
0.35615 | -73.648
0.35649 | -73.6467
0.35683 | -73.641
0.35717 | -73.6272
0.35751 | -73.6071
0.35785 | -73.5783
0.35819 | -73.5509
0.35853 | -73.5252
0.35887 | -73.5018
0.35922 | -73.4871
0.35956 | -73.4709
0.3599 | -73.4512
0.36024 | -73.4312
0.36058 | -73.4135
0.36092 | -73.4008
0.36126 | -73.3916
0.3616 | -73.3866
0.36194 | -73.3732
0.36228 | -73.354
0.36262 | -73.3388
0.36296 | -73.3281
0.3633 | -73.3364
0.36364 | -73.3254
0.36398 | -73.3225
0.36432 | -73.3202
0.36466 | -73.3231
0.365 | -73.3329
0.36534 | -73.3467
0.36568 | -73.3677
0.36602 | -73.3927
0.36636 | -73.4244
0.36671 | -73.4626
0.36705 | -73.4995
0.36739 | -73.5316
0.36773 | -73.569
0.36807 | -73.6081
0.36841 | -73.651
0.36875 | -73.7022
0.36909 | -73.762
0.36943 | -73.8304
0.36977 | -73.9012
0.37011 | -73.9689
0.37045 | -74.0389
0.37079 | -74.0962
0.37113 | -74.1589
0.37147 | -74.2468
0.37181 | -74.3213
0.37215 | -74.3856
0.37249 | -74.4442
0.37283 | -74.5107
0.37317 | -74.5587
0.37351 | -74.6278
0.37386 | -74.7011
0.3742 | -74.7651
0.37454 | -74.811
0.37488 | -74.8602
0.37522 | -74.917
0.37556 | -74.9608
0.3759 | -74.9862
0.37624 | -75.0073
0.37658 | -75.0253
0.37692 | -75.0314
0.37726 | -75.0372
0.3776 | -75.0498
0.37794 | -75.0537
0.37828 | -75.0485
0.37862 | -75.0538
0.37896 | -75.0477
0.3793 | -75.0299
0.37964 | -75.0084
0.37999 | -74.9932
0.38033 | -74.994
0.38067 | -75.0027
0.38101 | -74.9863
0.38135 | -74.952
0.38169 | -74.9123
0.38203 | -74.8734
0.38237 | -74.8149
0.38271 | -74.7581
0.38305 | -74.7208
0.38339 | -74.6822
0.38373 | -74.6481
0.38407 | -74.6085
0.38441 | -74.5668
0.38475 | -74.5227
0.38509 | -74.4824
0.38543 | -74.4326
0.38577 | -74.3671
0.38611 | -74.3
0.38645 | -74.2379
0.38679 | -74.1863
0.38714 | -74.1367
0.38748 | -74.0866
0.38782 | -74.0363
0.38816 | -73.9783
0.3885 | -73.9268
0.38884 | -73.8599
0.38918 | -73.786
0.38952 | -73.7213
0.38986 | -73.6591
0.3902 | -73.599
0.39054 | -73.5286
0.39088 | -73.4518
0.39122 | -73.3845
0.39156 | -73.3229
0.3919 | -73.2598
0.39224 | -73.194
0.39258 | -73.1312
0.39292 | -73.0768
0.39326 | -73.009
0.3936 | -72.9398
0.39394 | -72.8815
0.39428 | -72.8282
0.39463 | -72.7832
0.39497 | -72.7398
0.39531 | -72.6931
0.39565 | -72.6504
0.39599 | -72.6059
0.39633 | -72.5617
0.39667 | -72.5175
0.39701 | -72.48
0.39735 | -72.4475
0.39769 | -72.4181
0.39803 | -72.3964
0.39837 | -72.3786
0.39871 | -72.3554
0.39905 | -72.3388
0.39939 | -72.3312
0.39973 | -72.3297
0.40007 | -72.3256
0.40041 | -72.3276
0.40075 | -72.337
0.40109 | -72.3509
0.40143 | -72.3665
0.40178 | -72.3867
0.40212 | -72.4174
0.40246 | -72.483
0.4028 | -72.5184
0.40314 | -72.5561
0.40348 | -72.6033
0.40382 | -72.6553
0.40416 | -72.707
0.4045 | -72.7575
0.40484 | -72.8187
0.40518 | -72.8776
0.40552 | -72.9308
0.40586 | -72.9935
0.4062 | -73.0672
0.40654 | -73.1403
0.40688 | -73.2042
0.40722 | -73.2713
0.40756 | -73.3524
0.40791 | -73.4226
0.40825 | -73.4862
0.40859 | -73.5591
0.40893 | -73.6353
0.40927 | -73.7102
0.40961 | -73.7863
0.40995 | -73.8731
0.41029 | -73.9613
0.41063 | -74.0299
0.41097 | -74.0799
0.41131 | -74.1373
0.41165 | -74.1845
0.41199 | -74.225
0.41233 | -74.2337
0.41267 | -74.2728
0.41301 | -74.3105
0.41335 | -74.344
0.41369 | -74.367
0.41403 | -74.3916
0.41437 | -74.4033
0.41472 | -74.4199
0.41506 | -74.4312
0.4154 | -74.4203
0.41574 | -74.4025
0.41608 | -74.3861
0.41642 | -74.3709
0.41676 | -74.3596
0.4171 | -74.342
0.41744 | -74.3313
0.41778 | -74.3209
0.41812 | -74.296
0.41846 | -74.2646
0.4188 | -74.2381
0.41914 | -74.2202
0.41948 | -74.2045
0.41982 | -74.178
0.42016 | -74.1432
0.4205 | -74.1145
0.42084 | -74.0954
0.42118 | -74.0729
0.42152 | -74.0497
0.42186 | -74.0327
0.4222 | -74.051
0.42255 | -74.0562
0.42289 | -74.0648
0.42323 | -74.0646
0.42357 | -74.0579
0.42391 | -74.0457
0.42425 | -74.0499
0.42459 | -74.0571
0.42493 | -74.0669
0.42527 | -74.0819
0.42561 | -74.0997
0.42595 | -74.1239
0.42629 | -74.1455
0.42663 | -74.17
0.42697 | -74.1936
0.42731 | -74.2184
0.42765 | -74.2485
0.42799 | -74.2713
0.42833 | -74.2891
0.42867 | -74.3134
0.42901 | -74.339
0.42936 | -74.3693
0.4297 | -74.3974
0.43004 | -74.4257
0.43038 | -74.4534
0.43072 | -74.4735
0.43106 | -74.4922
0.4314 | -74.5048
0.43174 | -74.5112
0.43208 | -74.4998
0.43242 | -74.5004
0.43276 | -74.4927
0.4331 | -74.4929
0.43344 | -74.495
0.43378 | -74.4918
0.43412 | -74.4816
0.43446 | -74.4795
0.4348 | -74.4734
0.43514 | -74.471
0.43548 | -74.4686
0.43583 | -74.4481
0.43617 | -74.4247
0.43651 | -74.4055
0.43685 | -74.3896
0.43719 | -74.3726
0.43753 | -74.3414
0.43787 | -74.3187
0.43821 | -74.3025
0.43855 | -74.273
0.43889 | -74.2406
0.43923 | -74.2156
0.43957 | -74.1914
0.43991 | -74.1563
0.44025 | -74.1186
0.44059 | -74.0851
0.44093 | -74.0545
0.44127 | -74.0214
0.44161 | -73.9954
0.44195 | -73.9809
0.44229 | -73.9505
0.44264 | -73.9196
0.44298 | -73.893
0.44332 | -73.8674
0.44366 | -73.8413
0.444 | -73.8191
0.44434 | -73.8059
0.44468 | -73.7962
0.44502 | -73.8016
0.44536 | -73.7965
0.4457 | -73.7847
0.44604 | -73.7758
0.44638 | -73.7748
0.44672 | -73.7769
0.44706 | -73.78
0.4474 | -73.7857
0.44774 | -73.7962
0.44808 | -73.804
0.44842 | -73.8161
0.44876 | -73.8343
0.4491 | -73.8506
0.44944 | -73.8641
0.44978 | -73.8841
0.45012 | -73.9056
0.45047 | -73.9251
0.45081 | -73.9501
0.45115 | -73.9732
0.45149 | -73.9903
0.45183 | -74.0204
0.45217 | -74.0467
0.45251 | -74.0714
0.45285 | -74.0944
0.45319 | -74.1165
0.45353 | -74.1315
0.45387 | -74.152
0.45421 | -74.1756
0.45455 | -74.1999
0.45489 | -74.2308
0.45523 | -74.2702
0.45557 | -74.2965
0.45591 | -74.3179
0.45625 | -74.3311
0.45659 | -74.346
0.45693 | -74.3634
0.45728 | -74.3807
0.45762 | -74.397
0.45796 | -74.41
0.4583 | -74.4281
0.45864 | -74.4473
0.45898 | -74.4575
0.45932 | -74.4753
0.45966 | -74.494
0.46 | -74.5073
0.46034 | -74.5221
0.46068 | -74.5341
0.46102 | -74.5431
0.46136 | -74.552
0.4617 | -74.5708
0.46204 | -74.5819
0.46238 | -74.6011
0.46272 | -74.6258
0.46306 | -74.6467
0.4634 | -74.6669
0.46375 | -74.6913
0.46409 | -74.7117
0.46443 | -74.7364
0.46477 | -74.76
0.46511 | -74.7709
0.46545 | -74.7795
0.46579 | -74.793
0.46613 | -74.8179
0.46647 | -74.8333
0.46681 | -74.8427
0.46715 | -74.8536
0.46749 | -74.8571
0.46783 | -74.862
0.46817 | -74.8698
0.46851 | -74.8733
0.46885 | -74.8778
0.46919 | -74.8765
0.46953 | -74.8713
0.46987 | -74.8685
0.47021 | -74.8664
0.47056 | -74.862
0.4709 | -74.8634
0.47124 | -74.8537
0.47158 | -74.8479
0.47192 | -74.8426
0.47226 | -74.8456
0.4726 | -74.8472
0.47294 | -74.8466
0.47328 | -74.8411
0.47362 | -74.8291
0.47396 | -74.8209
0.4743 | -74.8122
0.47464 | -74.8052
0.47498 | -74.8031
0.47532 | -74.7985
0.47566 | -74.7867
0.476 | -74.7721
0.47634 | -74.763
0.47668 | -74.7512
0.47702 | -74.7356
0.47736 | -74.7212
0.4777 | -74.7084
0.47804 | -74.7035
0.47839 | -74.6977
0.47873 | -74.6965
0.47907 | -74.689
0.47941 | -74.678
0.47975 | -74.671
0.48009 | -74.6659
0.48043 | -74.659
0.48077 | -74.6573
0.48111 | -74.6734
0.48145 | -74.6763
0.48179 | -74.6766
0.48213 | -74.6795
0.48247 | -74.6834
0.48281 | -74.685
0.48315 | -74.6878
0.48349 | -74.7002
0.48383 | -74.7159
0.48417 | -74.7325
0.48451 | -74.7532
0.48485 | -74.7706
0.4852 | -74.784
0.48554 | -74.7938
0.48588 | -74.8087
0.48622 | -74.8265
0.48656 | -74.8484
0.4869 | -74.8704
0.48724 | -74.8821
0.48758 | -74.8954
0.48792 | -74.9103
0.48826 | -74.9214
0.4886 | -74.9349
0.48894 | -74.9493
0.48928 | -74.9582
0.48962 | -74.9747
0.48996 | -74.9937
0.4903 | -75.0113
0.49064 | -75.0265
0.49098 | -75.0365
0.49132 | -75.0529
0.49167 | -75.0691
0.49201 | -75.0841
0.49235 | -75.0981
0.49269 | -75.1136
0.49303 | -75.1285
0.49337 | -75.1428
0.49371 | -75.1482
0.49405 | -75.154
0.49439 | -75.1624
0.49473 | -75.1732
0.49507 | -75.1814
0.49541 | -75.1869
0.49575 | -75.1907
0.49609 | -75.187
0.49643 | -75.1871
0.49677 | -75.19
0.49711 | -75.1975
0.49745 | -75.2065
0.49779 | -75.2071
0.49813 | -75.2074
0.49848 | -75.2115
0.49882 | -75.2204
0.49916 | -75.2342
0.4995 | -75.2451
0.49984 | -75.2501
0.50018 | -75.2466
0.50052 | -75.2552
0.50086 | -75.2813
0.5012 | -75.2978
0.50154 | -75.3118
0.50188 | -75.3339
0.50222 | -75.3617
0.50256 | -75.3793
0.5029 | -75.3924
0.50324 | -75.4086
0.50358 | -75.4301
0.50392 | -75.456
0.50426 | -75.48
0.5046 | -75.5041
0.50494 | -75.5307
0.50528 | -75.5497
0.50562 | -75.571
0.50596 | -75.5955
0.50631 | -75.6216
0.50665 | -75.6472
0.50699 | -75.6702
0.50733 | -75.6994
0.50767 | -75.7285
0.50801 | -75.7605
0.50835 | -75.7961
0.50869 | -75.8257
0.50903 | -75.8536
0.50937 | -75.8835
0.50971 | -75.9134
0.51005 | -75.9381
0.51039 | -75.9663
0.51073 | -75.9964
0.51107 | -76.033
0.51141 | -76.0637
0.51175 | -76.0918
0.51209 | -76.1164
0.51243 | -76.1393
0.51277 | -76.1682
0.51312 | -76.1954
0.51346 | -76.2241
0.5138 | -76.2464
0.51414 | -76.2721
0.51448 | -76.3017
0.51482 | -76.3251
0.51516 | -76.3523
0.5155 | -76.3854
0.51584 | -76.4198
0.51618 | -76.439
0.51652 | -76.46
0.51686 | -76.4865
0.5172 | -76.5108
0.51754 | -76.5366
0.51788 | -76.5619
0.51822 | -76.5855
0.51856 | -76.6094
0.5189 | -76.6284
0.51924 | -76.6442
0.51958 | -76.663
0.51992 | -76.6788
0.52027 | -76.6933
0.52061 | -76.7137
0.52095 | -76.7362
0.52129 | -76.7702
0.52163 | -76.7991
0.52197 | -76.8168
0.52231 | -76.8316
0.52265 | -76.8542
0.52299 | -76.8854
0.52333 | -76.9095
0.52367 | -76.9292
0.52401 | -76.9446
0.52435 | -76.9527
0.52469 | -76.9703
0.52503 | -76.9949
0.52537 | -77.0129
0.52571 | -77.024
0.52605 | -77.0379
0.52639 | -77.052
0.52673 | -77.0618
0.52708 | -77.0742
0.52742 | -77.0858
0.52776 | -77.0973
0.5281 | -77.1121
0.52844 | -77.1258
0.52878 | -77.1397
0.52912 | -77.1586
0.52946 | -77.1764
0.5298 | -77.1987
0.53014 | -77.216
0.53048 | -77.2403
0.53082 | -77.2604
0.53116 | -77.2818
0.5315 | -77.3056
0.53184 | -77.3282
0.53218 | -77.3504
0.53252 | -77.3822
0.53286 | -77.4126
0.5332 | -77.4317
0.53355 | -77.45
0.53389 | -77.4672
0.53423 | -77.4937
0.53457 | -77.5233
0.53491 | -77.5442
0.53525 | -77.566
0.53559 | -77.5924
0.53593 | -77.6173
0.53627 | -77.6502
0.53661 | -77.6883
0.53695 | -77.7072
0.53729 | -77.7256
0.53763 | -77.7514
0.53797 | -77.7771
0.53831 | -77.8026
0.53865 | -77.8294
0.53899 | -77.8577
0.53933 | -77.8846
0.53967 | -77.9176
0.54001 | -77.9456
0.54036 | -77.9747
0.5407 | -77.9966
0.54104 | -78.0216
0.54138 | -78.0433
0.54172 | -78.0625
0.54206 | -78.0834
0.5424 | -78.107
0.54274 | -78.1268
0.54308 | -78.1483
0.54342 | -78.1712
0.54376 | -78.1906
0.5441 | -78.2076
0.54444 | -78.2136
0.54478 | -78.2178
0.54512 | -78.221
0.54546 | -78.2293
0.5458 | -78.2415
0.54614 | -78.2429
0.54648 | -78.2402
0.54682 | -78.244
0.54716 | -78.2504
0.54751 | -78.2564
0.54785 | -78.2562
0.54819 | -78.252
0.54853 | -78.2493
0.54887 | -78.2529
0.54921 | -78.2543
0.54955 | -78.2521
0.54989 | -78.2355
0.55023 | -78.2339
0.55057 | -78.2263
0.55091 | -78.2118
0.55125 | -78.204
0.55159 | -78.1995
0.55193 | -78.1821
0.55227 | -78.1646
0.55261 | -78.1562
0.55295 | -78.1425
0.55329 | -78.1297
0.55363 | -78.1224
0.55397 | -78.1226
0.55432 | -78.1157
0.55466 | -78.1044
0.555 | -78.1011
0.55534 | -78.1
0.55568 | -78.0959
0.55602 | -78.0858
0.55636 | -78.0769
0.5567 | -78.066
0.55704 | -78.0564
0.55738 | -78.0469
0.55772 | -78.0413
0.55806 | -78.0341
0.5584 | -78.0254
0.55874 | -78.0169
0.55908 | -78.0083
0.55942 | -77.9957
0.55976 | -77.9794
0.5601 | -77.9682
0.56044 | -77.9571
0.56078 | -77.9452
0.56112 | -77.9328
0.56146 | -77.925
0.5618 | -77.9198
0.56215 | -77.9155
0.56249 | -77.9222
0.56283 | -77.9316
0.56317 | -77.9286
0.56351 | -77.9204
0.56385 | -77.9088
0.56419 | -77.8993
0.56453 | -77.8878
0.56487 | -77.874
0.56521 | -77.8583
0.56555 | -77.8512
0.56589 | -77.8319
0.56623 | -77.8096
0.56657 | -77.7892
0.56691 | -77.7721
0.56725 | -77.756
0.56759 | -77.7373
0.56793 | -77.7185
0.56827 | -77.7014
0.56861 | -77.6876
0.56896 | -77.6687
0.5693 | -77.6434
0.56964 | -77.6108
0.56998 | -77.5911
0.57032 | -77.5685
0.57066 | -77.5476
0.571 | -77.5279
0.57134 | -77.5121
0.57168 | -77.4984
0.57202 | -77.4685
0.57236 | -77.4368
0.5727 | -77.4127
0.57304 | -77.3889
0.57338 | -77.3619
0.57372 | -77.336
0.57406 | -77.3162
0.5744 | -77.299
0.57474 | -77.2746
0.57508 | -77.2497
0.57542 | -77.2246
0.57576 | -77.2041
0.57611 | -77.1877
0.57645 | -77.1657
0.57679 | -77.1389
0.57713 | -77.1151
0.57747 | -77.0972
0.57781 | -77.0781
0.57815 | -77.0574
0.57849 | -77.0429
0.57883 | -77.0274
0.57917 | -77.0069
0.57951 | -77.0056
0.57985 | -76.9897
0.58019 | -76.9803
0.58053 | -76.9739
0.58087 | -76.9669
0.58121 | -76.9605
0.58155 | -76.9539
0.58189 | -76.9507
0.58223 | -76.9495
0.58257 | -76.953
0.58292 | -76.9551
0.58326 | -76.9562
0.5836 | -76.9623
0.58394 | -76.9709
0.58428 | -76.9852
0.58462 | -77.0019
0.58496 | -77.0201
0.5853 | -77.0361
0.58564 | -77.0484
0.58598 | -77.0624
0.58632 | -77.0851
0.58666 | -77.1091
0.587 | -77.1329
0.58734 | -77.1494
0.58768 | -77.168
0.58802 | -77.195
0.58836 | -77.2213
0.5887 | -77.2434
0.58904 | -77.2697
0.58939 | -77.3143
0.58973 | -77.3445
0.59007 | -77.3779
0.59041 | -77.4169
0.59075 | -77.4503
0.59109 | -77.487
0.59143 | -77.5284
0.59177 | -77.562
0.59211 | -77.5989
0.59245 | -77.6426
0.59279 | -77.6803
0.59313 | -77.71
0.59347 | -77.7461
0.59381 | -77.784
0.59415 | -77.8207
0.59449 | -77.8566
0.59483 | -77.8934
0.59517 | -77.9297
0.59551 | -77.9625
0.59585 | -77.998
0.5962 | -78.0341
0.59654 | -78.071
0.59688 | -78.1091
0.59722 | -78.1424
0.59756 | -78.1809
0.5979 | -78.214
0.59824 | -78.2459
0.59858 | -78.2787
0.59892 | -78.2821
0.59926 | -78.3083
0.5996 | -78.3195
0.59994 | -78.3311
0.60028 | -78.3563
0.60062 | -78.3853
0.60096 | -78.4064
0.6013 | -78.4273
0.60164 | -78.4507
0.60198 | -78.468
0.60232 | -78.4795
0.60266 | -78.4871
0.603 | -78.4893
0.60335 | -78.4937
0.60369 | -78.4981
0.60403 | -78.4995
0.60437 | -78.5022
0.60471 | -78.5085
0.60505 | -78.5177
0.60539 | -78.5238
0.60573 | -78.5273
0.60607 | -78.5343
0.60641 | -78.5421
0.60675 | -78.5363
0.60709 | -78.5418
0.60743 | -78.552
0.60777 | -78.5565
0.60811 | -78.5543
0.60845 | -78.5451
0.60879 | -78.5335
0.60913 | -78.5268
0.60947 | -78.5183
0.60981 | -78.5147
0.61016 | -78.5105
0.6105 | -78.5081
0.61084 | -78.5094
0.61118 | -78.5127
0.61152 | -78.5064
0.61186 | -78.5012
0.6122 | -78.5013
0.61254 | -78.4999
0.61288 | -78.4966
0.61322 | -78.4937
0.61356 | -78.4928
0.6139 | -78.4998
0.61424 | -78.5036
0.61458 | -78.5051
0.61492 | -78.5129
0.61526 | -78.5132
0.6156 | -78.5114
0.61594 | -78.5109
0.61628 | -78.5125
0.61662 | -78.5137
0.61696 | -78.5167
0.6173 | -78.5268
0.61764 | -78.5373
0.61799 | -78.5479
0.61833 | -78.5628
0.61867 | -78.5871
0.61901 | -78.6044
0.61935 | -78.62
0.61969 | -78.6317
0.62003 | -78.6484
0.62037 | -78.6642
0.62071 | -78.6758
0.62105 | -78.6947
0.62139 | -78.7184
0.62173 | -78.7376
0.62207 | -78.7582
0.62241 | -78.7809
0.62275 | -78.8039
0.62309 | -78.8235
0.62343 | -78.8356
0.62377 | -78.8543
0.62411 | -78.8792
0.62445 | -78.8992
0.6248 | -78.9207
0.62514 | -78.9408
0.62548 | -78.9635
0.62582 | -78.9893
0.62616 | -79.0146
0.6265 | -79.0398
0.62684 | -79.0648
0.62718 | -79.0875
0.62752 | -79.109
0.62786 | -79.1296
0.6282 | -79.1499
0.62854 | -79.1534
0.62888 | -79.1879
0.62922 | -79.2124
0.62956 | -79.2267
0.6299 | -79.2458
0.63024 | -79.2737
0.63058 | -79.2974
0.63092 | -79.3167
0.63126 | -79.3366
0.6316 | -79.3605
0.63195 | -79.3772
0.63229 | -79.3873
0.63263 | -79.3983
0.63297 | -79.4143
0.63331 | -79.4308
0.63365 | -79.4458
0.63399 | -79.4653
0.63433 | -79.4749
0.63467 | -79.4863
0.63501 | -79.5071
0.63535 | -79.5239
0.63569 | -79.5375
0.63603 | -79.5543
0.63637 | -79.5754
0.63671 | -79.584
0.63705 | -79.5888
0.63739 | -79.595
0.63773 | -79.6031
0.63807 | -79.6084
0.63841 | -79.6112
0.63876 | -79.6219
0.6391 | -79.6355
0.63944 | -79.6414
0.63978 | -79.6423
0.64012 | -79.6447
0.64046 | -79.6487
0.6408 | -79.661
0.64114 | -79.6653
0.64148 | -79.6681
0.64182 | -79.6717
0.64216 | -79.675
0.6425 | -79.6851
0.64284 | -79.698
0.64318 | -79.7039
0.64352 | -79.7185
0.64386 | -79.7326
0.6442 | -79.7317
0.64454 | -79.7238
0.64489 | -79.7187
0.64523 | -79.7143
0.64557 | -79.7102
0.64591 | -79.7083
0.64625 | -79.7111
0.64659 | -79.707
0.64693 | -79.7051
0.64727 | -79.7033
0.64761 | -79.7001
0.64795 | -79.6947
0.64829 | -79.6917
0.64863 | -79.6844
0.64897 | -79.6735
0.64931 | -79.6621
0.64965 | -79.6548
0.64999 | -79.6492
0.65033 | -79.6437
0.65067 | -79.6382
0.65101 | -79.6343
0.65135 | -79.6323
0.65169 | -79.6255
0.65204 | -79.62
0.65238 | -79.6196
0.65272 | -79.6203
0.65306 | -79.6223
0.6534 | -79.6204
0.65374 | -79.6211
0.65408 | -79.6211
0.65442 | -79.6124
0.65476 | -79.6047
0.6551 | -79.5985
0.65544 | -79.5953
0.65578 | -79.5909
0.65612 | -79.5904
0.65646 | -79.5894
0.6568 | -79.5856
0.65714 | -79.5835
0.65748 | -79.5838
0.65782 | -79.5844
0.65816 | -79.5888
0.6585 | -79.5905
0.65885 | -79.5972
0.65919 | -79.6041
0.65953 | -79.6076
0.65987 | -79.6106
0.66021 | -79.6189
0.66055 | -79.6255
0.66089 | -79.6343
0.66123 | -79.6417
0.66157 | -79.651
0.66191 | -79.6601
0.66225 | -79.6635
0.66259 | -79.6626
0.66293 | -79.6588
0.66327 | -79.6589
0.66361 | -79.6634
0.66395 | -79.6587
0.66429 | -79.6503
0.66463 | -79.6504
0.66497 | -79.6555
0.66531 | -79.6603
0.66565 | -79.6679
0.666 | -79.6796
0.66634 | -79.685
0.66668 | -79.6911
0.66702 | -79.7014
0.66736 | -79.7115
0.6677 | -79.7229
0.66804 | -79.7287
0.66838 | -79.7354
0.66872 | -79.7475
0.66906 | -79.7606
0.6694 | -79.7697
0.66974 | -79.7772
0.67008 | -79.7838
0.67042 | -79.793
0.67076 | -79.8015
0.6711 | -79.8089
0.67144 | -79.8118
0.67178 | -79.8163
0.67212 | -79.8248
0.67246 | -79.8312
0.6728 | -79.8414
0.67314 | -79.8484
0.67349 | -79.858
0.67383 | -79.8752
0.67417 | -79.8943
0.67451 | -79.9069
0.67485 | -79.9119
0.67519 | -79.9199
0.67553 | -79.9295
0.67587 | -79.9344
0.67621 | -79.9402
0.67655 | -79.9474
0.67689 | -79.9462
0.67723 | -79.9456
0.67757 | -79.9518
0.67791 | -79.9594
0.67825 | -79.9705
0.67859 | -79.9822
0.67893 | -79.9901
0.67927 | -79.9954
0.67961 | -79.9991
0.67995 | -80.0013
0.68029 | -80.001
0.68064 | -80.0093
0.68098 | -80.0173
0.68132 | -80.0248
0.68166 | -80.0306
0.682 | -80.0354
0.68234 | -80.0371
0.68268 | -80.0265
0.68302 | -80.0274
0.68336 | -80.0322
0.6837 | -80.0311
0.68404 | -80.0232
0.68438 | -80.0138
0.68472 | -80.0136
0.68506 | -80.0136
0.6854 | -80.0154
0.68574 | -80.0163
0.68608 | -80.0111
0.68642 | -80.0103
0.68676 | -80.012
0.6871 | -80.0133
0.68745 | -80.0186
0.68779 | -80.0211
0.68813 | -80.0243
0.68847 | -80.0306
0.68881 | -80.0332
0.68915 | -80.0324
0.68949 | -80.0333
0.68983 | -80.0356
0.69017 | -80.0379
0.69051 | -80.0466
0.69085 | -80.0518
0.69119 | -80.0567
0.69153 | -80.0673
0.69187 | -80.0766
0.69221 | -80.0823
0.69255 | -80.0867
0.69289 | -80.0868
0.69323 | -80.0873
0.69357 | -80.0935
0.69391 | -80.1017
0.69425 | -80.1086
0.6946 | -80.1157
0.69494 | -80.127
0.69528 | -80.1376
0.69562 | -80.1454
0.69596 | -80.1549
0.6963 | -80.1699
0.69664 | -80.1844
0.69698 | -80.1887
0.69732 | -80.1957
0.69766 | -80.2074
0.698 | -80.2154
0.69834 | -80.2227
0.69868 | -80.2319
0.69902 | -80.2413
0.69936 | -80.2472
0.6997 | -80.249
0.70004 | -80.249
0.70038 | -80.2532
0.70073 | -80.2614
0.70107 | -80.2694
0.70141 | -80.2772
0.70175 | -80.2859
0.70209 | -80.2964
0.70243 | -80.3064
0.70277 | -80.3148
0.70311 | -80.3256
0.70345 | -80.3345
0.70379 | -80.338
0.70413 | -80.3482
0.70447 | -80.3606
0.70481 | -80.3713
0.70515 | -80.3787
0.70549 | -80.3858
0.70583 | -80.3998
0.70617 | -80.4152
0.70651 | -80.425
0.70685 | -80.4273
0.70719 | -80.4394
0.70753 | -80.4561
0.70788 | -80.467
0.70822 | -80.4767
0.70856 | -80.4858
0.7089 | -80.4947
0.70924 | -80.5053
0.70958 | -80.5184
0.70992 | -80.5311
0.71026 | -80.5456
0.7106 | -80.563
0.71094 | -80.5826
0.71128 | -80.6076
0.71162 | -80.6258
0.71196 | -80.6382
0.7123 | -80.6574
0.71264 | -80.6784
0.71298 | -80.6885
0.71332 | -80.6942
0.71366 | -80.7055
0.714 | -80.715
0.71434 | -80.7214
0.71469 | -80.7317
0.71503 | -80.7371
0.71537 | -80.7381
0.71571 | -80.7408
0.71605 | -80.747
0.71639 | -80.7571
0.71673 | -80.7645
0.71707 | -80.7645
0.71741 | -80.7721
0.71775 | -80.7785
0.71809 | -80.7891
0.71843 | -80.7963
0.71877 | -80.8028
0.71911 | -80.809
0.71945 | -80.8089
0.71979 | -80.8152
0.72013 | -80.8204
0.72047 | -80.8332
0.72081 | -80.8407
0.72115 | -80.8444
0.72149 | -80.8487
0.72184 | -80.851
0.72218 | -80.8555
0.72252 | -80.8596
0.72286 | -80.8616
0.7232 | -80.8625
0.72354 | -80.8655
0.72388 | -80.8704
0.72422 | -80.8721
0.72456 | -80.8739
0.7249 | -80.8766
0.72524 | -80.8784
0.72558 | -80.8825
0.72592 | -80.8896
0.72626 | -80.8991
0.7266 | -80.9015
0.72694 | -80.9072
0.72728 | -80.9091
0.72762 | -80.9133
0.72796 | -80.9187
0.7283 | -80.9247
0.72864 | -80.9307
0.72898 | -80.9372
0.72933 | -80.9461
0.72967 | -80.9563
0.73001 | -80.9709
0.73035 | -80.9779
0.73069 | -80.9881
0.73103 | -80.9999
0.73137 | -81.0136
0.73171 | -81.0273
0.73205 | -81.0422
0.73239 | -81.077
0.73273 | -81.1002
0.73307 | -81.1166
0.73341 | -81.1351
0.73375 | -81.1622
0.73409 | -81.196
0.73443 | -81.2185
0.73477 | -81.2205
0.73511 | -81.2238
0.73545 | -81.234
0.73579 | -81.2575
0.73613 | -81.278
0.73648 | -81.3066
0.73682 | -81.3296
0.73716 | -81.3476
0.7375 | -81.3649
0.73784 | -81.3848
0.73818 | -81.4057
0.73852 | -81.427
0.73886 | -81.4504
0.7392 | -81.4731
0.73954 | -81.5018
0.73988 | -81.528
0.74022 | -81.5518
0.74056 | -81.5772
0.7409 | -81.601
0.74124 | -81.6241
0.74158 | -81.6473
0.74192 | -81.6683
0.74226 | -81.6878
0.7426 | -81.7092
0.74294 | -81.7334
0.74329 | -81.7639
0.74363 | -81.7939
0.74397 | -81.8137
0.74431 | -81.8361
0.74465 | -81.8568
0.74499 | -81.8718
0.74533 | -81.8937
0.74567 | -81.9168
0.74601 | -81.94
0.74635 | -81.9559
0.74669 | -81.981
0.74703 | -82.0012
0.74737 | -82.0238
0.74771 | -82.0524
0.74805 | -82.079
0.74839 | -82.0951
0.74873 | -82.1001
0.74907 | -82.1168
0.74941 | -82.1337
0.74975 | -82.1491
0.75009 | -82.1638
0.75044 | -82.1838
0.75078 | -82.2052
0.75112 | -82.2258
0.75146 | -82.2454
0.7518 | -82.2499
0.75214 | -82.2548
0.75248 | -82.2608
0.75282 | -82.2676
0.75316 | -82.2775
0.7535 | -82.2839
0.75384 | -82.2889
0.75418 | -82.292
0.75452 | -82.2997
0.75486 | -82.3129
0.7552 | -82.3117
0.75554 | -82.3052
0.75588 | -82.3034
0.75622 | -82.3049
0.75657 | -82.3121
0.75691 | -82.3218
0.75725 | -82.3222
0.75759 | -82.3177
0.75793 | -82.3236
0.75827 | -82.3422
0.75861 | -82.3499
0.75895 | -82.3492
0.75929 | -82.3484
0.75963 | -82.3478
0.75997 | -82.3463
0.76031 | -82.3487
0.76065 | -82.3485
0.76099 | -82.3346
0.76133 | -82.3292
0.76167 | -82.3294
0.76201 | -82.3293
0.76235 | -82.3248
0.76269 | -82.3226
0.76303 | -82.3228
0.76337 | -82.3267
0.76372 | -82.3322
0.76406 | -82.342
0.7644 | -82.3521
0.76474 | -82.3619
0.76508 | -82.3741
0.76542 | -82.3782
0.76576 | -82.3789
0.7661 | -82.375
0.76644 | -82.3792
0.76678 | -82.3822
0.76712 | -82.388
0.76746 | -82.3925
0.7678 | -82.3879
0.76814 | -82.3782
0.76848 | -82.3714
0.76882 | -82.3748
0.76916 | -82.3824
0.7695 | -82.3812
0.76984 | -82.3789
0.77018 | -82.378
0.77053 | -82.3778
0.77087 | -82.3802
0.77121 | -82.3811
0.77155 | -82.3847
0.77189 | -82.3911
0.77223 | -82.3933
0.77257 | -82.3952
0.77291 | -82.3959
0.77325 | -82.3985
0.77359 | -82.3963
0.77393 | -82.3847
0.77427 | -82.3825
0.77461 | -82.3857
0.77495 | -82.3788
0.77529 | -82.3736
0.77563 | -82.3718
0.77597 | -82.3648
0.77631 | -82.3617
0.77665 | -82.3588
0.77699 | -82.3568
0.77733 | -82.3504
0.77768 | -82.3445
0.77802 | -82.3399
0.77836 | -82.3361
0.7787 | -82.3352
0.77904 | -82.332
0.77938 | -82.3277
0.77972 | -82.3278
0.78006 | -82.3277
0.7804 | -82.323
0.78074 | -82.3154
0.78108 | -82.3089
0.78142 | -82.3021
0.78176 | -82.2965
0.7821 | -82.2932
0.78244 | -82.2883
0.78278 | -82.2847
0.78312 | -82.2774
0.78346 | -82.2686
0.7838 | -82.2637
0.78414 | -82.2597
0.78448 | -82.2542
0.78482 | -82.2508
0.78517 | -82.2518
0.78551 | -82.2484
0.78585 | -82.2458
0.78619 | -82.2423
0.78653 | -82.2361
0.78687 | -82.2327
0.78721 | -82.2294
0.78755 | -82.2267
0.78789 | -82.2185
0.78823 | -82.2219
0.78857 | -82.2244
0.78891 | -82.2231
0.78925 | -82.221
0.78959 | -82.2223
0.78993 | -82.2215
0.79027 | -82.2177
0.79061 | -82.2142
0.79095 | -82.2076
0.79129 | -82.203
0.79163 | -82.2
0.79197 | -82.1983
0.79232 | -82.1972
0.79266 | -82.1985
0.793 | -82.2044
0.79334 | -82.2003
0.79368 | -82.1961
0.79402 | -82.1972
0.79436 | -82.1983
0.7947 | -82.2004
0.79504 | -82.202
0.79538 | -82.2042
0.79572 | -82.2132
0.79606 | -82.2184
0.7964 | -82.2272
0.79674 | -82.2361
0.79708 | -82.2424
0.79742 | -82.2476
0.79776 | -82.2512
0.7981 | -82.2612
0.79844 | -82.2686
0.79878 | -82.2779
0.79913 | -82.2935
0.79947 | -82.3006
0.79981 | -82.3066
0.80015 | -82.3146
0.80049 | -82.3218
0.80083 | -82.3307
0.80117 | -82.3386
0.80151 | -82.3483
0.80185 | -82.3576
0.80219 | -82.3693
0.80253 | -82.3816
0.80287 | -82.395
0.80321 | -82.4168
0.80355 | -82.428
0.80389 | -82.4352
0.80423 | -82.4421
0.80457 | -82.4525
0.80491 | -82.4635
0.80525 | -82.4806
0.80559 | -82.4919
0.80593 | -82.5062
0.80628 | -82.5185
0.80662 | -82.5327
0.80696 | -82.5469
0.8073 | -82.5603
0.80764 | -82.5728
0.80798 | -82.5825
0.80832 | -82.5964
0.80866 | -82.6088
0.809 | -82.6182
0.80934 | -82.6298
0.80968 | -82.6437
0.81002 | -82.6631
0.81036 | -82.6868
0.8107 | -82.7059
0.81104 | -82.7152
0.81138 | -82.7209
0.81172 | -82.7345
0.81206 | -82.7498
0.81241 | -82.764
0.81275 | -82.7814
0.81309 | -82.7977
0.81343 | -82.8107
0.81377 | -82.8253
0.81411 | -82.8385
0.81445 | -82.852
0.81479 | -82.8662
0.81513 | -82.8801
0.81547 | -82.888
0.81581 | -82.8806
0.81615 | -82.8903
0.81649 | -82.9183
0.81683 | -82.9197
0.81717 | -82.9219
0.81751 | -82.9303
0.81785 | -82.944
0.81819 | -82.9537
0.81853 | -82.9596
0.81887 | -82.9657
0.81921 | -82.9711
0.81956 | -82.9779
0.8199 | -82.9832
0.82024 | -82.9906
0.82058 | -82.997
0.82092 | -83.0051
0.82126 | -83.0149
0.8216 | -83.0283
0.82194 | -83.0376
0.82228 | -83.04
0.82262 | -83.0394
0.82296 | -83.0399
0.8233 | -83.0458
0.82364 | -83.0581
0.82398 | -83.0693
0.82432 | -83.0691
0.82466 | -83.063
0.825 | -83.0503
0.82534 | -83.0441
0.82568 | -83.0364
0.82602 | -83.0334
0.82637 | -83.0312
0.82671 | -83.0279
0.82705 | -83.0231
0.82739 | -83.0154
0.82773 | -83.0041
0.82807 | -82.9939
0.82841 | -82.9847
0.82875 | -82.9732
0.82909 | -82.9653
0.82943 | -82.9568
0.82977 | -82.9498
0.83011 | -82.953
0.83045 | -82.9385
0.83079 | -82.9224
0.83113 | -82.9205
0.83147 | -82.9137
0.83181 | -82.9054
0.83215 | -82.8955
0.83249 | -82.8853
0.83283 | -82.8749
0.83317 | -82.8614
0.83352 | -82.8473
0.83386 | -82.8403
0.8342 | -82.8323
0.83454 | -82.8187
0.83488 | -82.8049
0.83522 | -82.7921
0.83556 | -82.7792
0.8359 | -82.7679
0.83624 | -82.7567
0.83658 | -82.7441
0.83692 | -82.7294
0.83726 | -82.7172
0.8376 | -82.7096
0.83794 | -82.6931
0.83828 | -82.6784
0.83862 | -82.669
0.83896 | -82.6569
0.8393 | -82.6474
0.83964 | -82.6385
0.83998 | -82.6323
0.84032 | -82.6312
0.84066 | -82.6205
0.84101 | -82.6099
0.84135 | -82.6027
0.84169 | -82.5935
0.84203 | -82.5849
0.84237 | -82.5795
0.84271 | -82.579
0.84305 | -82.5752
0.84339 | -82.573
0.84373 | -82.5713
0.84407 | -82.5688
0.84441 | -82.569
0.84475 | -82.5766
0.84509 | -82.5733
0.84543 | -82.5732
0.84577 | -82.578
0.84611 | -82.584
0.84645 | -82.5854
0.84679 | -82.5855
0.84713 | -82.5824
0.84747 | -82.5764
0.84781 | -82.575
0.84816 | -82.5769
0.8485 | -82.5842
0.84884 | -82.594
0.84918 | -82.5995
0.84952 | -82.6033
0.84986 | -82.6092
0.8502 | -82.6154
0.85054 | -82.6211
0.85088 | -82.6339
0.85122 | -82.6458
0.85156 | -82.6539
0.8519 | -82.6651
0.85224 | -82.6769
0.85258 | -82.6861
0.85292 | -82.6974
0.85326 | -82.7079
0.8536 | -82.7197
0.85394 | -82.7328
0.85428 | -82.7433
0.85462 | -82.7644
0.85497 | -82.7769
0.85531 | -82.7899
0.85565 | -82.8027
0.85599 | -82.8183
0.85633 | -82.8328
0.85667 | -82.8487
0.85701 | -82.8652
0.85735 | -82.8821
0.85769 | -82.8992
0.85803 | -82.9136
0.85837 | -82.9288
0.85871 | -82.9471
0.85905 | -82.9629
0.85939 | -82.9804
0.85973 | -82.9986
0.86007 | -83.0128
0.86041 | -83.0278
0.86075 | -83.0453
0.86109 | -83.0651
0.86143 | -83.0823
0.86177 | -83.1071
0.86212 | -83.133
0.86246 | -83.155
0.8628 | -83.1734
0.86314 | -83.1908
0.86348 | -83.2072
0.86382 | -83.2256
0.86416 | -83.2361
0.8645 | -83.2478
0.86484 | -83.2609
0.86518 | -83.2778
0.86552 | -83.2986
0.86586 | -83.3154
0.8662 | -83.3333
0.86654 | -83.3533
0.86688 | -83.3705
0.86722 | -83.3756
0.86756 | -83.3837
0.8679 | -83.3912
0.86825 | -83.3939
0.86859 | -83.405
0.86893 | -83.4176
0.86927 | -83.428
0.86961 | -83.4354
0.86995 | -83.4517
0.87029 | -83.4696
0.87063 | -83.4764
0.87097 | -83.48
0.87131 | -83.4853
0.87165 | -83.485
0.87199 | -83.4864
0.87233 | -83.4911
0.87267 | -83.4966
0.87301 | -83.5035
0.87335 | -83.5119
0.87369 | -83.5215
0.87403 | -83.5068
0.87437 | -83.5073
0.87471 | -83.506
0.87505 | -83.517
0.8754 | -83.5346
0.87574 | -83.5399
0.87608 | -83.5369
0.87642 | -83.5444
0.87676 | -83.5454
0.8771 | -83.5408
0.87744 | -83.5346
0.87778 | -83.5299
0.87812 | -83.526
0.87846 | -83.5179
0.8788 | -83.5118
0.87914 | -83.5047
0.87948 | -83.4988
0.87982 | -83.484
0.88016 | -83.4695
0.8805 | -83.4584
0.88084 | -83.4434
0.88118 | -83.4279
0.88152 | -83.4164
0.88186 | -83.4042
0.88221 | -83.3932
0.88255 | -83.385
0.88289 | -83.3662
0.88323 | -83.3477
0.88357 | -83.3344
0.88391 | -83.3118
0.88425 | -83.3028
0.88459 | -83.2895
0.88493 | -83.2738
0.88527 | -83.2588
0.88561 | -83.244
0.88595 | -83.2298
0.88629 | -83.2091
0.88663 | -83.1856
0.88697 | -83.1642
0.88731 | -83.1436
0.88765 | -83.1246
0.88799 | -83.1115
0.88833 | -83.0966
0.88867 | -83.0788
0.88901 | -83.0635
0.88936 | -83.0453
0.8897 | -83.0202
0.89004 | -82.9947
0.89038 | -82.9747
0.89072 | -82.9603
0.89106 | -82.9452
0.8914 | -82.9211
0.89174 | -82.8887
0.89208 | -82.8631
0.89242 | -82.8434
0.89276 | -82.8227
0.8931 | -82.8018
0.89344 | -82.7813
0.89378 | -82.7627
0.89412 | -82.7451
0.89446 | -82.7277
0.8948 | -82.7099
0.89514 | -82.6938
0.89548 | -82.6781
0.89582 | -82.6623
0.89616 | -82.6463
0.8965 | -82.6288
0.89685 | -82.6098
0.89719 | -82.5951
0.89753 | -82.579
0.89787 | -82.5645
0.89821 | -82.5524
0.89855 | -82.5381
0.89889 | -82.5303
0.89923 | -82.5212
0.89957 | -82.509
0.89991 | -82.4971
0.90025 | -82.4807
0.90059 | -82.4687
0.90093 | -82.4509
0.90127 | -82.4398
0.90161 | -82.4333
0.90195 | -82.4256
0.90229 | -82.4168
0.90263 | -82.409
0.90297 | -82.4003
0.90331 | -82.3919
0.90365 | -82.3949
0.904 | -82.3862
0.90434 | -82.3756
0.90468 | -82.365
0.90502 | -82.3625
0.90536 | -82.3605
0.9057 | -82.3556
0.90604 | -82.3515
0.90638 | -82.3503
0.90672 | -82.3461
0.90706 | -82.338
0.9074 | -82.3256
0.90774 | -82.3193
0.90808 | -82.3128
0.90842 | -82.3052
0.90876 | -82.3041
0.9091 | -82.3052
0.90944 | -82.3033
0.90978 | -82.2996
0.91012 | -82.2963
0.91046 | -82.2926
0.91081 | -82.2836
0.91115 | -82.2793
0.91149 | -82.2769
0.91183 | -82.2759
0.91217 | -82.2744
0.91251 | -82.2703
0.91285 | -82.2666
0.91319 | -82.2615
0.91353 | -82.2645
0.91387 | -82.2613
0.91421 | -82.2591
0.91455 | -82.2591
0.91489 | -82.2563
0.91523 | -82.2559
0.91557 | -82.2557
0.91591 | -82.2519
0.91625 | -82.2492
0.91659 | -82.2483
0.91694 | -82.2476
0.91727 | -82.2458
0.91761 | -82.2439
0.91796 | -82.2449
0.9183 | -82.2443
0.91864 | -82.2449
0.91898 | -82.2444
0.91932 | -82.2445
0.91966 | -82.2403
0.92 | -82.2319
0.92034 | -82.2299
0.92068 | -82.2266
0.92102 | -82.2208
0.92136 | -82.2146
0.9217 | -82.2111
0.92204 | -82.2098
0.92238 | -82.2104
0.92272 | -82.2095
0.92306 | -82.2053
0.9234 | -82.1929
0.92374 | -82.1847
0.92409 | -82.1791
0.92443 | -82.1717
0.92477 | -82.1665
0.92511 | -82.163
0.92545 | -82.1601
0.92579 | -82.1572
0.92613 | -82.1538
0.92647 | -82.1492
0.92681 | -82.1497
0.92715 | -82.1456
0.92749 | -82.134
0.92783 | -82.1259
0.92817 | -82.1198
0.92851 | -82.1127
0.92885 | -82.1074
0.92919 | -82.0986
0.92953 | -82.0905
0.92987 | -82.0859
0.93021 | -82.0868
0.93055 | -82.0881
0.9309 | -82.0872
0.93124 | -82.0865
0.93158 | -82.0797
0.93192 | -82.0705
0.93226 | -82.0642
0.9326 | -82.0604
0.93294 | -82.0577
0.93328 | -82.0481
0.93362 | -82.0361
0.93396 | -82.0277
0.9343 | -82.0209
0.93464 | -82.014
0.93498 | -82.0095
0.93532 | -82.0015
0.93566 | -81.9946
0.936 | -81.9836
0.93634 | -81.9736
0.93668 | -81.9664
0.93702 | -81.9571
0.93736 | -81.9486
0.9377 | -81.947
0.93805 | -81.9441
0.93839 | -81.938
0.93873 | -81.9309
0.93907 | -81.9289
0.93941 | -81.9257
0.93975 | -81.921
0.94009 | -81.916
0.94043 | -81.9095
0.94077 | -81.9036
0.94111 | -81.8987
0.94145 | -81.8941
0.94179 | -81.8898
0.94213 | -81.8845
0.94247 | -81.8791
0.94281 | -81.8846
0.94315 | -81.8827
0.94349 | -81.8816
0.94383 | -81.8847
0.94417 | -81.8867
0.94451 | -81.8858
0.94486 | -81.8899
0.9452 | -81.8905
0.94554 | -81.8886
0.94588 | -81.8905
0.94622 | -81.8935
0.94656 | -81.8938
0.9469 | -81.8978
0.94724 | -81.9014
0.94758 | -81.9039
0.94792 | -81.9054
0.94826 | -81.8965
0.9486 | -81.8899
0.94894 | -81.8898
0.94928 | -81.8934
0.94962 | -81.897
0.94996 | -81.8943
0.9503 | -81.8897
0.95064 | -81.8918
0.95098 | -81.8968
0.95132 | -81.9011
0.95166 | -81.9051
0.952 | -81.9103
0.95234 | -81.9144
0.95269 | -81.9254
0.95303 | -81.9277
0.95337 | -81.9281
0.95371 | -81.9281
0.95405 | -81.9308
0.95439 | -81.9389
0.95473 | -81.949
0.95507 | -81.9545
0.95541 | -81.959
0.95575 | -81.965
0.95609 | -81.9719
0.95643 | -81.9843
0.95677 | -81.9912
0.95711 | -81.9959
0.95745 | -82.0011
0.95779 | -82.0071
0.95813 | -82.0122
0.95847 | -82.0161
0.95881 | -82.0234
0.95915 | -82.0298
0.9595 | -82.0338
0.95984 | -82.0372
0.96018 | -82.0431
0.96052 | -82.0497
0.96086 | -82.055
0.9612 | -82.0591
0.96154 | -82.0664
0.96188 | -82.0755
0.96222 | -82.0807
0.96256 | -82.0825
0.9629 | -82.0879
0.96324 | -82.0972
0.96358 | -82.1022
0.96392 | -82.1058
0.96426 | -82.1086
0.9646 | -82.1142
0.96494 | -82.1187
0.96528 | -82.1228
0.96562 | -82.1311
0.96596 | -82.1318
0.9663 | -82.1289
0.96665 | -82.1285
0.96699 | -82.1323
0.96733 | -82.1366
0.96767 | -82.1407
0.96801 | -82.1435
0.96835 | -82.1457
0.96869 | -82.1431
0.96903 | -82.1413
0.96937 | -82.1419
0.96971 | -82.1427
0.97005 | -82.1434
0.97039 | -82.1449
0.97073 | -82.1428
0.97107 | -82.1403
0.97141 | -82.1364
0.97175 | -82.1324
0.97209 | -82.1306
0.97243 | -82.1282
0.97277 | -82.1297
0.97311 | -82.1383
0.97346 | -82.1481
0.9738 | -82.1501
0.97414 | -82.1521
0.97448 | -82.1509
0.97482 | -82.1476
0.97516 | -82.1456
0.9755 | -82.1436
0.97584 | -82.1414
0.97618 | -82.1397
0.97652 | -82.1363
0.97686 | -82.1332
0.9772 | -82.1259
0.97754 | -82.1183
0.97788 | -82.1107
0.97822 | -82.1027
0.97856 | -82.0965
0.9789 | -82.0903
0.97924 | -82.0837
0.97958 | -82.0751
0.97993 | -82.0672
0.98027 | -82.0621
0.98061 | -82.0537
0.98095 | -82.0452
0.98129 | -82.0393
0.98163 | -82.0298
0.98197 | -82.0206
0.98231 | -82.014
0.98265 | -82.0098
0.98299 | -82.0056
0.98333 | -81.9912
0.98367 | -81.9721
0.98401 | -81.9608
0.98435 | -81.9504
0.98469 | -81.9395
0.98503 | -81.9305
0.98537 | -81.9208
0.98571 | -81.9118
0.98605 | -81.9012
0.98639 | -81.892
0.98674 | -81.8889
0.98708 | -81.883
0.98742 | -81.8749
0.98776 | -81.8665
0.9881 | -81.8605
0.98844 | -81.8536
0.98878 | -81.8447
0.98912 | -81.8361
0.98946 | -81.8272
0.9898 | -81.8179
0.99014 | -81.8107
0.99048 | -81.8046
0.99082 | -81.7947
0.99116 | -81.7806
0.9915 | -81.772
0.99184 | -81.7619
0.99218 | -81.7561
0.99252 | -81.7454
0.99286 | -81.7332
0.9932 | -81.7217
0.99354 | -81.7122
0.99389 | -81.7034
0.99423 | -81.6947
0.99457 | -81.6882
0.99491 | -81.6852
0.99525 | -81.6783
0.99559 | -81.667
0.99593 | -81.6577
0.99627 | -81.6485
0.99661 | -81.6411
0.99695 | -81.6298
0.99729 | -81.618
0.99763 | -81.6075
0.99797 | -81.5983
0.99831 | -81.5916
0.99865 | -81.5858
0.99899 | -81.5791
0.99933 | -81.5723
0.99967 | -81.5624
1 | -81.5512
1.0004 | -81.5429
1.0007 | -81.5369
1.001 | -81.5336
1.0014 | -81.5337
1.0017 | -81.5265
1.0021 | -81.5165
1.0024 | -81.5093
1.0027 | -81.5038
1.0031 | -81.5003
1.0034 | -81.4962
1.0038 | -81.4894
1.0041 | -81.4866
1.0044 | -81.4838
1.0048 | -81.4778
1.0051 | -81.4723
1.0055 | -81.4688
1.0058 | -81.4646
1.0061 | -81.4592
1.0065 | -81.4529
1.0068 | -81.4499
1.0072 | -81.4454
1.0075 | -81.438
1.0078 | -81.4315
1.0082 | -81.4268
1.0085 | -81.4244
1.0089 | -81.4225
1.0092 | -81.4194
1.0095 | -81.4176
1.0099 | -81.417
1.0102 | -81.4168
1.0106 | -81.4173
1.0109 | -81.4164
1.0112 | -81.4159
1.0116 | -81.42
1.0119 | -81.419
1.0123 | -81.4186
1.0126 | -81.4196
1.013 | -81.4187
1.0133 | -81.4167
1.0136 | -81.4175
1.014 | -81.4212
1.0143 | -81.4242
1.0147 | -81.4244
1.015 | -81.4248
1.0153 | -81.427
1.0157 | -81.4306
1.016 | -81.4348
1.0164 | -81.4392
1.0167 | -81.4458
1.017 | -81.4573
1.0174 | -81.4668
1.0177 | -81.4732
1.0181 | -81.4749
1.0184 | -81.4755
1.0187 | -81.4829
1.0191 | -81.4927
1.0194 | -81.4982
1.0198 | -81.5011
1.0201 | -81.5047
1.0204 | -81.5087
1.0208 | -81.5098
1.0211 | -81.512
1.0215 | -81.5227
1.0218 | -81.5264
1.0221 | -81.5275
1.0225 | -81.5328
1.0228 | -81.5432
1.0232 | -81.5539
1.0235 | -81.562
1.0238 | -81.5698
1.0242 | -81.5764
1.0245 | -81.5809
1.0249 | -81.5881
1.0252 | -81.5948
1.0255 | -81.603
1.0259 | -81.6112
1.0262 | -81.6191
1.0266 | -81.6258
1.0269 | -81.6325
1.0273 | -81.6386
1.0276 | -81.645
1.0279 | -81.6518
1.0283 | -81.6596
1.0286 | -81.6677
1.029 | -81.6762
1.0293 | -81.6874
1.0296 | -81.6979
1.03 | -81.708
1.0303 | -81.7196
1.0307 | -81.7311
1.031 | -81.7418
1.0313 | -81.7549
1.0317 | -81.7639
1.032 | -81.7746
1.0324 | -81.7876
1.0327 | -81.8027
1.033 | -81.8138
1.0334 | -81.822
1.0337 | -81.8351
1.0341 | -81.85
1.0344 | -81.8644
1.0347 | -81.8772
1.0351 | -81.889
1.0354 | -81.9031
1.0358 | -81.9174
1.0361 | -81.9352
1.0364 | -81.9607
1.0368 | -81.9771
1.0371 | -81.9853
1.0375 | -81.9974
1.0378 | -82.01
1.0381 | -82.0249
1.0385 | -82.0418
1.0388 | -82.0592
1.0392 | -82.0702
1.0395 | -82.0839
1.0399 | -82.1007
1.0402 | -82.1167
1.0405 | -82.1278
1.0409 | -82.1397
1.0412 | -82.1504
1.0416 | -82.165
1.0419 | -82.1826
1.0422 | -82.1989
1.0426 | -82.2124
1.0429 | -82.2244
1.0433 | -82.237
1.0436 | -82.2483
1.0439 | -82.2597
1.0443 | -82.2732
1.0446 | -82.2866
1.045 | -82.3016
1.0453 | -82.3194
1.0456 | -82.3343
1.046 | -82.348
1.0463 | -82.3622
1.0467 | -82.377
1.047 | -82.3872
1.0473 | -82.3926
1.0477 | -82.4043
1.048 | -82.4211
1.0484 | -82.4359
1.0487 | -82.4492
1.049 | -82.4636
1.0494 | -82.4782
1.0497 | -82.493
1.0501 | -82.5075
1.0504 | -82.5184
1.0507 | -82.53
1.0511 | -82.5413
1.0514 | -82.5526
1.0518 | -82.5672
1.0521 | -82.5817
1.0524 | -82.5936
1.0528 | -82.6024
1.0531 | -82.6088
1.0535 | -82.6176
1.0538 | -82.6274
1.0542 | -82.633
1.0545 | -82.6393
1.0548 | -82.6524
1.0552 | -82.6656
1.0555 | -82.6706
1.0559 | -82.674
1.0562 | -82.6768
1.0565 | -82.6779
1.0569 | -82.6797
1.0572 | -82.6845
1.0576 | -82.6906
1.0579 | -82.6958
1.0582 | -82.6996
1.0586 | -82.7034
1.0589 | -82.708
1.0593 | -82.7107
1.0596 | -82.7105
1.0599 | -82.7064
1.0603 | -82.6993
1.0606 | -82.6971
1.061 | -82.6843
1.0613 | -82.6738
1.0616 | -82.6707
1.062 | -82.6613
1.0623 | -82.6435
1.0627 | -82.636
1.063 | -82.6343
1.0633 | -82.6241
1.0637 | -82.6147
1.064 | -82.6075
1.0644 | -82.6003
1.0647 | -82.5941
1.065 | -82.5895
1.0654 | -82.584
1.0657 | -82.5754
1.0661 | -82.5644
1.0664 | -82.5528
1.0667 | -82.5397
1.0671 | -82.527
1.0674 | -82.5193
1.0678 | -82.5125
1.0681 | -82.5002
1.0685 | -82.4823
1.0688 | -82.4676
1.0691 | -82.4539
1.0695 | -82.4418
1.0698 | -82.4308
1.0702 | -82.4159
1.0705 | -82.3942
1.0708 | -82.3785
1.0712 | -82.361
1.0715 | -82.3406
1.0719 | -82.3213
1.0722 | -82.3072
1.0725 | -82.2934
1.0729 | -82.2751
1.0732 | -82.2561
1.0736 | -82.2381
1.0739 | -82.2168
1.0742 | -82.1984
1.0746 | -82.1868
1.0749 | -82.1752
1.0753 | -82.1583
1.0756 | -82.1397
1.0759 | -82.119
1.0763 | -82.0965
1.0766 | -82.0747
1.077 | -82.0535
1.0773 | -82.0341
1.0776 | -82.0147
1.078 | -81.9911
1.0783 | -81.9677
1.0787 | -81.947
1.079 | -81.926
1.0793 | -81.9059
1.0797 | -81.8858
1.08 | -81.8665
1.0804 | -81.8491
1.0807 | -81.8331
1.0811 | -81.816
1.0814 | -81.7971
1.0817 | -81.7726
1.0821 | -81.7468
1.0824 | -81.7259
1.0828 | -81.7064
1.0831 | -81.6856
1.0834 | -81.6625
1.0838 | -81.6384
1.0841 | -81.6181
1.0845 | -81.5985
1.0848 | -81.5809
1.0851 | -81.5633
1.0855 | -81.5455
1.0858 | -81.5284
1.0862 | -81.5104
1.0865 | -81.4936
1.0868 | -81.4793
1.0872 | -81.4621
1.0875 | -81.4446
1.0879 | -81.4292
1.0882 | -81.4145
1.0885 | -81.3998
1.0889 | -81.3856
1.0892 | -81.3722
1.0896 | -81.3589
1.0899 | -81.3454
1.0902 | -81.3367
1.0906 | -81.3205
1.0909 | -81.309
1.0913 | -81.2978
1.0916 | -81.2854
1.0919 | -81.2771
1.0923 | -81.2679
1.0926 | -81.2554
1.093 | -81.2455
1.0933 | -81.237
1.0936 | -81.2248
1.094 | -81.2142
1.0943 | -81.2067
1.0947 | -81.2014
1.095 | -81.1961
1.0954 | -81.1902
1.0957 | -81.1822
1.096 | -81.1746
1.0964 | -81.1704
1.0967 | -81.1667
1.0971 | -81.1605
1.0974 | -81.154
1.0977 | -81.1494
1.0981 | -81.1445
1.0984 | -81.1402
1.0988 | -81.1378
1.0991 | -81.1353
1.0994 | -81.132
1.0998 | -81.1281
1.1001 | -81.1369
1.1005 | -81.1368
1.1008 | -81.1381
1.1011 | -81.1395
1.1015 | -81.1419
1.1018 | -81.1453
1.1022 | -81.1491
1.1025 | -81.1531
1.1028 | -81.1536
1.1032 | -81.1535
1.1035 | -81.1559
1.1039 | -81.157
1.1042 | -81.1604
1.1045 | -81.1639
1.1049 | -81.1661
1.1052 | -81.1686
1.1056 | -81.1701
1.1059 | -81.1729
1.1062 | -81.1797
1.1066 | -81.1865
1.1069 | -81.1923
1.1073 | -81.1993
1.1076 | -81.2056
1.1079 | -81.2141
1.1083 | -81.2227
1.1086 | -81.2276
1.109 | -81.2356
1.1093 | -81.2428
1.1097 | -81.2503
1.11 | -81.2623
1.1103 | -81.2712
1.1107 | -81.2805
1.111 | -81.2868
1.1114 | -81.293
1.1117 | -81.3016
1.112 | -81.3128
1.1124 | -81.3251
1.1127 | -81.3314
1.1131 | -81.3415
1.1134 | -81.3521
1.1137 | -81.3616
1.1141 | -81.3717
1.1144 | -81.3872
1.1148 | -81.3989
1.1151 | -81.4078
1.1154 | -81.4201
1.1158 | -81.4362
1.1161 | -81.4503
1.1165 | -81.4629
1.1168 | -81.4749
1.1171 | -81.4843
1.1175 | -81.4911
1.1178 | -81.5009
1.1182 | -81.5118
1.1185 | -81.5209
1.1188 | -81.5287
1.1192 | -81.5383
1.1195 | -81.5497
1.1199 | -81.5609
1.1202 | -81.5717
1.1205 | -81.5828
1.1209 | -81.5937
1.1212 | -81.6034
1.1216 | -81.6125
1.1219 | -81.6233
1.1222 | -81.6351
1.1226 | -81.6471
1.1229 | -81.662
1.1233 | -81.6821
1.1236 | -81.6984
1.124 | -81.7094
1.1243 | -81.7162
1.1246 | -81.7242
1.125 | -81.7336
1.1253 | -81.7427
1.1257 | -81.7519
1.126 | -81.7615
1.1263 | -81.7721
1.1267 | -81.7815
1.127 | -81.7874
1.1274 | -81.7957
1.1277 | -81.8046
1.128 | -81.8099
1.1284 | -81.8165
1.1287 | -81.8261
1.1291 | -81.8389
1.1294 | -81.852
1.1297 | -81.8615
1.1301 | -81.871
1.1304 | -81.8794
1.1308 | -81.8905
1.1311 | -81.897
1.1314 | -81.9066
1.1318 | -81.9203
1.1321 | -81.9318
1.1325 | -81.94
1.1328 | -81.95
1.1331 | -81.9618
1.1335 | -81.9719
1.1338 | -81.9787
1.1342 | -81.986
1.1345 | -81.9957
1.1348 | -82.0042
1.1352 | -82.0111
1.1355 | -82.018
1.1359 | -82.0264
1.1362 | -82.0362
1.1365 | -82.0476
1.1369 | -82.0595
1.1372 | -82.0657
1.1376 | -82.072
1.1379 | -82.0836
1.1383 | -82.0975
1.1386 | -82.1059
1.1389 | -82.1198
1.1393 | -82.1179
1.1396 | -82.1233
1.14 | -82.1333
1.1403 | -82.1418
1.1406 | -82.1494
1.141 | -82.1553
1.1413 | -82.1591
1.1417 | -82.1575
1.142 | -82.1524
1.1423 | -82.1479
1.1427 | -82.1491
1.143 | -82.153
1.1434 | -82.1575
1.1437 | -82.1619
1.144 | -82.1682
1.1444 | -82.172
1.1447 | -82.1736
1.1451 | -82.1753
1.1454 | -82.1798
1.1457 | -82.1869
1.1461 | -82.1909
1.1464 | -82.1958
1.1468 | -82.1972
1.1471 | -82.1967
1.1474 | -82.1998
1.1478 | -82.2029
1.1481 | -82.2047
1.1485 | -82.2056
1.1488 | -82.2084
1.1491 | -82.212
1.1495 | -82.2179
1.1498 | -82.2239
1.1502 | -82.2274
1.1505 | -82.2311
1.1508 | -82.2345
1.1512 | -82.2378
1.1515 | -82.2401
1.1519 | -82.2439
1.1522 | -82.2439
1.1526 | -82.2454
1.1529 | -82.2487
1.1532 | -82.2513
1.1536 | -82.2505
1.1539 | -82.2496
1.1543 | -82.2521
1.1546 | -82.2561
1.1549 | -82.2585
1.1553 | -82.2587
1.1556 | -82.2601
1.156 | -82.2683
1.1563 | -82.2754
1.1566 | -82.2774
1.157 | -82.2794
1.1573 | -82.2832
1.1577 | -82.2858
1.158 | -82.2877
1.1583 | -82.2902
1.1587 | -82.2914
1.159 | -82.2956
1.1594 | -82.2964
1.1597 | -82.2977
1.16 | -82.3021
1.1604 | -82.3112
1.1607 | -82.3198
1.1611 | -82.3246
1.1614 | -82.3245
1.1617 | -82.3263
1.1621 | -82.3326
1.1624 | -82.3362
1.1628 | -82.3371
1.1631 | -82.3379
1.1634 | -82.3399
1.1638 | -82.3429
1.1641 | -82.3435
1.1645 | -82.3436
1.1648 | -82.347
1.1651 | -82.3522
1.1655 | -82.3574
1.1658 | -82.3624
1.1662 | -82.3683
1.1665 | -82.3727
1.1669 | -82.3751
1.1672 | -82.3787
1.1675 | -82.3828
1.1679 | -82.3879
1.1682 | -82.3927
1.1686 | -82.3982
1.1689 | -82.4082
1.1692 | -82.414
1.1696 | -82.4192
1.1699 | -82.4258
1.1703 | -82.4292
1.1706 | -82.4323
1.1709 | -82.4378
1.1713 | -82.4437
1.1716 | -82.4503
1.172 | -82.4577
1.1723 | -82.4622
1.1726 | -82.4665
1.173 | -82.477
1.1733 | -82.4835
1.1737 | -82.4881
1.174 | -82.4951
1.1743 | -82.5002
1.1747 | -82.5059
1.175 | -82.5151
1.1754 | -82.526
1.1757 | -82.5347
1.176 | -82.5408
1.1764 | -82.5486
1.1767 | -82.5579
1.1771 | -82.5612
1.1774 | -82.5663
1.1777 | -82.576
1.1781 | -82.5863
1.1784 | -82.5946
1.1788 | -82.6027
1.1791 | -82.6103
1.1795 | -82.6205
1.1798 | -82.6288
1.1801 | -82.6358
1.1805 | -82.6438
1.1808 | -82.6509
1.1812 | -82.6584
1.1815 | -82.6655
1.1818 | -82.6702
1.1822 | -82.6777
1.1825 | -82.6883
1.1829 | -82.6999
1.1832 | -82.7104
1.1835 | -82.721
1.1839 | -82.7325
1.1842 | -82.7411
1.1846 | -82.7493
1.1849 | -82.7588
1.1852 | -82.7682
1.1856 | -82.78
1.1859 | -82.7905
1.1863 | -82.7985
1.1866 | -82.8069
1.1869 | -82.8147
1.1873 | -82.8207
1.1876 | -82.8298
1.188 | -82.8382
1.1883 | -82.8444
1.1886 | -82.8519
1.189 | -82.8607
1.1893 | -82.8671
1.1897 | -82.8776
1.19 | -82.887
1.1903 | -82.8961
1.1907 | -82.9077
1.191 | -82.914
1.1914 | -82.9214
1.1917 | -82.9331
1.192 | -82.9451
1.1924 | -82.9549
1.1927 | -82.9632
1.1931 | -82.9723
1.1934 | -82.9867
1.1938 | -82.9965
1.1941 | -83.006
1.1944 | -83.0139
1.1948 | -83.0213
1.1951 | -83.0301
1.1955 | -83.0398
1.1958 | -83.0507
1.1961 | -83.0591
1.1965 | -83.0664
1.1968 | -83.0754
1.1972 | -83.0836
1.1975 | -83.0939
1.1978 | -83.1071
1.1982 | -83.1183
1.1985 | -83.1328
1.1989 | -83.1444
1.1992 | -83.1553
1.1995 | -83.1671
1.1999 | -83.1816
1.2002 | -83.1925
1.2006 | -83.2
1.2009 | -83.2068
1.2012 | -83.2134
1.2016 | -83.221
1.2019 | -83.2296
1.2023 | -83.2362
1.2026 | -83.2411
1.2029 | -83.2504
1.2033 | -83.2594
1.2036 | -83.2679
1.204 | -83.2767
1.2043 | -83.2839
1.2046 | -83.2903
1.205 | -83.2988
1.2053 | -83.3052
1.2057 | -83.3172
1.206 | -83.3291
1.2063 | -83.3381
1.2067 | -83.3476
1.207 | -83.3567
1.2074 | -83.3667
1.2077 | -83.3765
1.2081 | -83.3856
1.2084 | -83.3953
1.2087 | -83.4077
1.2091 | -83.4162
1.2094 | -83.425
1.2098 | -83.4387
1.2101 | -83.4506
1.2104 | -83.4552
1.2108 | -83.4617
1.2111 | -83.4707
1.2115 | -83.4778
1.2118 | -83.4857
1.2121 | -83.4915
1.2125 | -83.4967
1.2128 | -83.5049
1.2132 | -83.5171
1.2135 | -83.5248
1.2138 | -83.53
1.2142 | -83.5363
1.2145 | -83.5475
1.2149 | -83.5574
1.2152 | -83.5636
1.2155 | -83.5704
1.2159 | -83.5778
1.2162 | -83.5823
1.2166 | -83.5879
1.2169 | -83.5946
1.2172 | -83.6021
1.2176 | -83.6107
1.2179 | -83.6188
1.2183 | -83.6263
1.2186 | -83.6307
1.2189 | -83.6346
1.2193 | -83.6422
1.2196 | -83.653
1.22 | -83.6611
1.2203 | -83.6664
1.2207 | -83.6721
1.221 | -83.6801
1.2213 | -83.6867
1.2217 | -83.6926
1.222 | -83.7107
1.2224 | -83.7338
1.2227 | -83.7405
1.223 | -83.7448
1.2234 | -83.7469
1.2237 | -83.7496
1.2241 | -83.7545
1.2244 | -83.7641
1.2247 | -83.7738
1.2251 | -83.7842
1.2254 | -83.7889
1.2258 | -83.7931
1.2261 | -83.798
1.2264 | -83.8037
1.2268 | -83.8068
1.2271 | -83.8117
1.2275 | -83.8173
1.2278 | -83.8231
1.2281 | -83.8289
1.2285 | -83.833
1.2288 | -83.8358
1.2292 | -83.8406
1.2295 | -83.8475
1.2298 | -83.8553
1.2302 | -83.8612
1.2305 | -83.8635
1.2309 | -83.8674
1.2312 | -83.8733
1.2315 | -83.8795
1.2319 | -83.882
1.2322 | -83.8859
1.2326 | -83.8911
1.2329 | -83.8953
1.2332 | -83.8979
1.2336 | -83.9008
1.2339 | -83.8985
1.2343 | -83.8995
1.2346 | -83.8995
1.2349 | -83.9035
1.2353 | -83.905
1.2356 | -83.9107
1.236 | -83.9131
1.2363 | -83.9142
1.2367 | -83.9164
1.237 | -83.9247
1.2373 | -83.9346
1.2377 | -83.9393
1.238 | -83.9469
1.2384 | -83.9498
1.2387 | -83.9533
1.239 | -83.9582
1.2394 | -83.9619
1.2397 | -83.9618
1.2401 | -83.9635
1.2404 | -83.9657
1.2407 | -83.9689
1.2411 | -83.9717
1.2414 | -83.9709
1.2418 | -83.9709
1.2421 | -83.9721
1.2424 | -83.9721
1.2428 | -83.9743
1.2431 | -83.9747
1.2435 | -83.976
1.2438 | -83.9752
1.2441 | -83.9717
1.2445 | -83.9717
1.2448 | -83.9742
1.2452 | -83.975
1.2455 | -83.9779
1.2458 | -83.9801
1.2462 | -83.9837
1.2465 | -83.9908
1.2469 | -83.9923
1.2472 | -83.9899
1.2475 | -83.9828
1.2479 | -83.9839
1.2482 | -83.9828
1.2486 | -83.9862
1.2489 | -83.9879
1.2493 | -83.9881
1.2496 | -83.9866
1.2499 | -83.9822
1.2503 | -83.9762
1.2506 | -83.9704
1.251 | -83.9674
1.2513 | -83.9668
1.2516 | -83.9672
1.252 | -83.9663
1.2523 | -83.9646
1.2527 | -83.9612
1.253 | -83.9557
1.2533 | -83.9506
1.2537 | -83.9467
1.254 | -83.9448
1.2544 | -83.9451
1.2547 | -83.9445
1.255 | -83.9447
1.2554 | -83.9468
1.2557 | -83.9427
1.2561 | -83.9352
1.2564 | -83.9287
1.2567 | -83.9247
1.2571 | -83.924
1.2574 | -83.9164
1.2578 | -83.9088
1.2581 | -83.906
1.2584 | -83.9038
1.2588 | -83.8979
1.2591 | -83.8917
1.2595 | -83.8863
1.2598 | -83.8812
1.2601 | -83.871
1.2605 | -83.8639
1.2608 | -83.8599
1.2612 | -83.856
1.2615 | -83.8516
1.2618 | -83.8479
1.2622 | -83.848
1.2625 | -83.847
1.2629 | -83.8406
1.2632 | -83.8324
1.2636 | -83.8243
1.2639 | -83.8197
1.2642 | -83.8161
1.2646 | -83.8115
1.2649 | -83.8058
1.2653 | -83.8002
1.2656 | -83.794
1.2659 | -83.7866
1.2663 | -83.7809
1.2666 | -83.774
1.267 | -83.7686
1.2673 | -83.7647
1.2676 | -83.7627
1.268 | -83.7589
1.2683 | -83.7529
1.2687 | -83.748
1.269 | -83.743
1.2693 | -83.7369
1.2697 | -83.731
1.27 | -83.7242
1.2704 | -83.7184
1.2707 | -83.7151
1.271 | -83.7133
1.2714 | -83.7091
1.2717 | -83.7034
1.2721 | -83.6974
1.2724 | -83.6913
1.2727 | -83.6885
1.2731 | -83.6843
1.2734 | -83.6759
1.2738 | -83.6698
1.2741 | -83.666
1.2744 | -83.662
1.2748 | -83.6582
1.2751 | -83.6557
1.2755 | -83.6521
1.2758 | -83.6461
1.2761 | -83.6401
1.2765 | -83.6347
1.2768 | -83.6345
1.2772 | -83.6298
1.2775 | -83.6283
1.2779 | -83.6248
1.2782 | -83.6178
1.2785 | -83.6061
1.2789 | -83.5973
1.2792 | -83.593
1.2796 | -83.5889
1.2799 | -83.5832
1.2802 | -83.577
1.2806 | -83.5712
1.2809 | -83.5657
1.2813 | -83.563
1.2816 | -83.5576
1.2819 | -83.5528
1.2823 | -83.5495
1.2826 | -83.5479
1.283 | -83.5461
1.2833 | -83.5435
1.2836 | -83.5434
1.284 | -83.5419
1.2843 | -83.5389
1.2847 | -83.5336
1.285 | -83.526
1.2853 | -83.517
1.2857 | -83.5105
1.286 | -83.5094
1.2864 | -83.5094
1.2867 | -83.5106
1.287 | -83.5054
1.2874 | -83.5015
1.2877 | -83.5001
1.2881 | -83.5017
1.2884 | -83.5039
1.2887 | -83.4988
1.2891 | -83.4912
1.2894 | -83.4898
1.2898 | -83.492
1.2901 | -83.4906
1.2905 | -83.4902
1.2908 | -83.4875
1.2911 | -83.4859
1.2915 | -83.4864
1.2918 | -83.4876
1.2922 | -83.4868
1.2925 | -83.4907
1.2928 | -83.4948
1.2932 | -83.4975
1.2935 | -83.4996
1.2939 | -83.5023
1.2942 | -83.5031
1.2945 | -83.5032
1.2949 | -83.5041
1.2952 | -83.5061
1.2956 | -83.5083
1.2959 | -83.5111
1.2962 | -83.5138
1.2966 | -83.5185
1.2969 | -83.5216
1.2973 | -83.5242
1.2976 | -83.5256
1.2979 | -83.5262
1.2983 | -83.5301
1.2986 | -83.5325
1.299 | -83.5349
1.2993 | -83.536
1.2996 | -83.5359
1.3 | -83.538
1.3003 | -83.5411
1.3007 | -83.5379
1.301 | -83.5403
1.3013 | -83.5443
1.3017 | -83.547
1.302 | -83.5508
1.3024 | -83.5534
1.3027 | -83.5551
1.303 | -83.5581
1.3034 | -83.5616
1.3037 | -83.5643
1.3041 | -83.5704
1.3044 | -83.575
1.3048 | -83.5788
1.3051 | -83.5814
1.3054 | -83.5821
1.3058 | -83.5838
1.3061 | -83.5885
1.3065 | -83.5894
1.3068 | -83.5912
1.3071 | -83.5964
1.3075 | -83.6023
1.3078 | -83.6067
1.3082 | -83.6101
1.3085 | -83.6144
1.3088 | -83.6197
1.3092 | -83.6229
1.3095 | -83.6244
1.3099 | -83.6274
1.3102 | -83.6317
1.3105 | -83.631
1.3109 | -83.6297
1.3112 | -83.6308
1.3116 | -83.6336
1.3119 | -83.6399
1.3122 | -83.6435
1.3126 | -83.6432
1.3129 | -83.6444
1.3133 | -83.6511
1.3136 | -83.6508
1.3139 | -83.651
1.3143 | -83.6518
1.3146 | -83.6503
1.315 | -83.6496
1.3153 | -83.6502
1.3156 | -83.648
1.316 | -83.6447
1.3163 | -83.6448
1.3167 | -83.6459
1.317 | -83.6465
1.3173 | -83.6478
1.3177 | -83.6493
1.318 | -83.6478
1.3184 | -83.6429
1.3187 | -83.6395
1.3191 | -83.6395
1.3194 | -83.6398
1.3197 | -83.6407
1.3201 | -83.6398
1.3204 | -83.6391
1.3208 | -83.6383
1.3211 | -83.638
1.3214 | -83.6369
1.3218 | -83.6362
1.3221 | -83.6357
1.3225 | -83.636
1.3228 | -83.6361
1.3231 | -83.6404
1.3235 | -83.6431
1.3238 | -83.6434
1.3242 | -83.6378
1.3245 | -83.6333
1.3248 | -83.6327
1.3252 | -83.6301
1.3255 | -83.6281
1.3259 | -83.6276
1.3262 | -83.6207
1.3265 | -83.6186
1.3269 | -83.612
1.3272 | -83.603
1.3276 | -83.5965
1.3279 | -83.5901
1.3282 | -83.5874
1.3286 | -83.585
1.3289 | -83.581
1.3293 | -83.5738
1.3296 | -83.5694
1.3299 | -83.5662
1.3303 | -83.5629
1.3306 | -83.5605
1.331 | -83.5548
1.3313 | -83.5503
1.3316 | -83.5473
1.332 | -83.5448
1.3323 | -83.5443
1.3327 | -83.5411
1.333 | -83.5374
1.3334 | -83.5333
1.3337 | -83.5329
1.334 | -83.532
1.3344 | -83.5262
1.3347 | -83.5191
1.3351 | -83.5156
1.3354 | -83.513
1.3357 | -83.5088
1.3361 | -83.5035
1.3364 | -83.4986
1.3368 | -83.4951
1.3371 | -83.4925
1.3374 | -83.4919
1.3378 | -83.4902
1.3381 | -83.4856
1.3385 | -83.4808
1.3388 | -83.4771
1.3391 | -83.4722
1.3395 | -83.4696
1.3398 | -83.466
1.3402 | -83.4631
1.3405 | -83.4605
1.3408 | -83.4587
1.3412 | -83.4586
1.3415 | -83.4553
1.3419 | -83.4538
1.3422 | -83.4498
1.3425 | -83.4458
1.3429 | -83.4424
1.3432 | -83.4409
1.3436 | -83.4405
1.3439 | -83.4402
1.3442 | -83.4392
1.3446 | -83.4383
1.3449 | -83.438
1.3453 | -83.4364
1.3456 | -83.4389
1.3459 | -83.4372
1.3463 | -83.435
1.3466 | -83.4347
1.347 | -83.4361
1.3473 | -83.4353
1.3477 | -83.4358
1.348 | -83.4342
1.3483 | -83.4311
1.3487 | -83.4331
1.349 | -83.4367
1.3494 | -83.4335
1.3497 | -83.4329
1.35 | -83.4332
1.3504 | -83.4297
1.3507 | -83.4262
1.3511 | -83.4258
1.3514 | -83.4269
1.3517 | -83.4234
1.3521 | -83.4197
1.3524 | -83.4167
1.3528 | -83.4153
1.3531 | -83.4135
1.3534 | -83.4115
1.3538 | -83.4115
1.3541 | -83.4118
1.3545 | -83.4131
1.3548 | -83.4154
1.3551 | -83.4153
1.3555 | -83.4158
1.3558 | -83.4119
1.3562 | -83.4108
1.3565 | -83.4112
1.3568 | -83.412
1.3572 | -83.4131
1.3575 | -83.4145
1.3579 | -83.4141
1.3582 | -83.416
1.3585 | -83.4187
1.3589 | -83.4196
1.3592 | -83.4187
1.3596 | -83.4179
1.3599 | -83.42
1.3602 | -83.4216
1.3606 | -83.4232
1.3609 | -83.4266
1.3613 | -83.4298
1.3616 | -83.4329
1.362 | -83.4364
1.3623 | -83.4396
1.3626 | -83.4414
1.363 | -83.439
1.3633 | -83.437
1.3637 | -83.4415
1.364 | -83.446
1.3643 | -83.45
1.3647 | -83.4518
1.365 | -83.4519
1.3654 | -83.4556
1.3657 | -83.4621
1.366 | -83.4679
1.3664 | -83.4727
1.3667 | -83.4756
1.3671 | -83.4788
1.3674 | -83.4834
1.3677 | -83.4879
1.3681 | -83.4902
1.3684 | -83.4905
1.3688 | -83.4928
1.3691 | -83.4972
1.3694 | -83.5017
1.3698 | -83.5057
1.3701 | -83.5105
1.3705 | -83.5145
1.3708 | -83.5201
1.3711 | -83.5312
1.3715 | -83.5387
1.3718 | -83.5437
1.3722 | -83.5478
1.3725 | -83.5518
1.3728 | -83.5557
1.3732 | -83.5571
1.3735 | -83.5645
1.3739 | -83.5692
1.3742 | -83.5748
1.3746 | -83.5807
1.3749 | -83.5848
1.3752 | -83.5875
1.3756 | -83.5902
1.3759 | -83.594
1.3763 | -83.5985
1.3766 | -83.6031
1.3769 | -83.6061
1.3773 | -83.6093
1.3776 | -83.617
1.378 | -83.6266
1.3783 | -83.6366
1.3786 | -83.6434
1.379 | -83.6463
1.3793 | -83.6507
1.3797 | -83.6568
1.38 | -83.6607
1.3803 | -83.6645
1.3807 | -83.6707
1.381 | -83.6766
1.3814 | -83.681
1.3817 | -83.6823
1.382 | -83.6888
1.3824 | -83.6985
1.3827 | -83.704
1.3831 | -83.7089
1.3834 | -83.7162
1.3837 | -83.7201
1.3841 | -83.7236
1.3844 | -83.7261
1.3848 | -83.729
1.3851 | -83.7312
1.3854 | -83.7339
1.3858 | -83.7388
1.3861 | -83.743
1.3865 | -83.7451
1.3868 | -83.7471
1.3871 | -83.7511
1.3875 | -83.7536
1.3878 | -83.7561
1.3882 | -83.7567
1.3885 | -83.7581
1.3889 | -83.761
1.3892 | -83.764
1.3895 | -83.7668
1.3899 | -83.7698
1.3902 | -83.7741
1.3906 | -83.779
1.3909 | -83.782
1.3912 | -83.7849
1.3916 | -83.7889
1.3919 | -83.7943
1.3923 | -83.7997
1.3926 | -83.8032
1.3929 | -83.8049
1.3933 | -83.8078
1.3936 | -83.813
1.394 | -83.8186
1.3943 | -83.8247
1.3946 | -83.8344
1.395 | -83.8331
1.3953 | -83.8295
1.3957 | -83.8274
1.396 | -83.8308
1.3963 | -83.8343
1.3967 | -83.8358
1.397 | -83.8374
1.3974 | -83.8393
1.3977 | -83.8399
1.398 | -83.8417
1.3984 | -83.8427
1.3987 | -83.8446
1.3991 | -83.8473
1.3994 | -83.8498
1.3997 | -83.8526
1.4001 | -83.8546
1.4004 | -83.8517
1.4008 | -83.8505
1.4011 | -83.849
1.4014 | -83.8521
1.4018 | -83.8544
1.4021 | -83.8557
1.4025 | -83.8536
1.4028 | -83.8515
1.4032 | -83.8523
1.4035 | -83.8543
1.4038 | -83.8568
1.4042 | -83.8587
1.4045 | -83.8561
1.4049 | -83.8516
1.4052 | -83.8529
1.4055 | -83.8577
1.4059 | -83.8615
1.4062 | -83.8651
1.4066 | -83.8667
1.4069 | -83.865
1.4072 | -83.865
1.4076 | -83.8699
1.4079 | -83.8718
1.4083 | -83.8708
1.4086 | -83.8735
1.4089 | -83.8766
1.4093 | -83.8786
1.4096 | -83.8802
1.41 | -83.8782
1.4103 | -83.878
1.4106 | -83.8822
1.411 | -83.8851
1.4113 | -83.8849
1.4117 | -83.8859
1.412 | -83.8849
1.4123 | -83.8847
1.4127 | -83.8859
1.413 | -83.8869
1.4134 | -83.8916
1.4137 | -83.8988
1.414 | -83.9021
1.4144 | -83.9066
1.4147 | -83.908
1.4151 | -83.9101
1.4154 | -83.9126
1.4158 | -83.915
1.4161 | -83.9168
1.4164 | -83.9198
1.4168 | -83.9236
1.4171 | -83.9267
1.4175 | -83.929
1.4178 | -83.9306
1.4181 | -83.9305
1.4185 | -83.9311
1.4188 | -83.9317
1.4192 | -83.9351
1.4195 | -83.9398
1.4198 | -83.9462
1.4202 | -83.9514
1.4205 | -83.9541
1.4209 | -83.9548
1.4212 | -83.9582
1.4215 | -83.9601
1.4219 | -83.9632
1.4222 | -83.9684
1.4226 | -83.9703
1.4229 | -83.9768
1.4232 | -83.9848
1.4236 | -83.9899
1.4239 | -83.9953
1.4243 | -84.0004
1.4246 | -84.0037
1.4249 | -84.0042
1.4253 | -84
1.4256 | -84.0025
1.426 | -84.0068
1.4263 | -84.0099
1.4266 | -84.0131
1.427 | -84.0157
1.4273 | -84.0221
1.4277 | -84.0271
1.428 | -84.0301
1.4283 | -84.0319
1.4287 | -84.0293
1.429 | -84.0298
1.4294 | -84.0332
1.4297 | -84.0365
1.4301 | -84.0364
1.4304 | -84.0368
1.4307 | -84.0381
1.4311 | -84.035
1.4314 | -84.0297
1.4318 | -84.0298
1.4321 | -84.0322
1.4324 | -84.0352
1.4328 | -84.0383
1.4331 | -84.0419
1.4335 | -84.0425
1.4338 | -84.0407
1.4341 | -84.0386
1.4345 | -84.0351
1.4348 | -84.0299
1.4352 | -84.029
1.4355 | -84.0273
1.4358 | -84.0236
1.4362 | -84.0214
1.4365 | -84.0211
1.4369 | -84.021
1.4372 | -84.0206
1.4375 | -84.0211
1.4379 | -84.0235
1.4382 | -84.0233
1.4386 | -84.0249
1.4389 | -84.0237
1.4392 | -84.0245
1.4396 | -84.0242
1.4399 | -84.0213
1.4403 | -84.0195
1.4406 | -84.0191
1.4409 | -84.0187
1.4413 | -84.018
1.4416 | -84.0155
1.442 | -84.0139
1.4423 | -84.0122
1.4426 | -84.0109
1.443 | -84.0098
1.4433 | -84.0081
1.4437 | -84.0034
1.444 | -83.9925
1.4444 | -83.9913
1.4447 | -83.9854
1.445 | -83.9789
1.4454 | -83.9723
1.4457 | -83.97
1.4461 | -83.9693
1.4464 | -83.966
1.4467 | -83.9632
1.4471 | -83.963
1.4474 | -83.9589
1.4478 | -83.9551
1.4481 | -83.9564
1.4484 | -83.9565
1.4488 | -83.9517
1.4491 | -83.9506
1.4495 | -83.9518
1.4498 | -83.9518
1.4501 | -83.9501
1.4505 | -83.945
1.4508 | -83.943
1.4512 | -83.9387
1.4515 | -83.929
1.4518 | -83.9224
1.4522 | -83.9192
1.4525 | -83.9157
1.4529 | -83.913
1.4532 | -83.9113
1.4535 | -83.9061
1.4539 | -83.8939
1.4542 | -83.8923
1.4546 | -83.8887
1.4549 | -83.884
1.4552 | -83.8781
1.4556 | -83.8724
1.4559 | -83.8685
1.4563 | -83.8649
1.4566 | -83.8596
1.4569 | -83.8533
1.4573 | -83.8482
1.4576 | -83.8447
1.458 | -83.8408
1.4583 | -83.8329
1.4587 | -83.8258
1.459 | -83.8209
1.4593 | -83.8162
1.4597 | -83.8111
1.46 | -83.8053
1.4604 | -83.8001
1.4607 | -83.7957
1.461 | -83.7908
1.4614 | -83.7892
1.4617 | -83.7876
1.4621 | -83.7821
1.4624 | -83.7765
1.4627 | -83.768
1.4631 | -83.759
1.4634 | -83.7547
1.4638 | -83.7435
1.4641 | -83.7351
1.4644 | -83.7285
1.4648 | -83.7195
1.4651 | -83.712
1.4655 | -83.7061
1.4658 | -83.6993
1.4661 | -83.692
1.4665 | -83.6817
1.4668 | -83.671
1.4672 | -83.6635
1.4675 | -83.6592
1.4678 | -83.6541
1.4682 | -83.6456
1.4685 | -83.6362
1.4689 | -83.628
1.4692 | -83.6219
1.4695 | -83.6172
1.4699 | -83.611
1.4702 | -83.6054
1.4706 | -83.5992
1.4709 | -83.5899
1.4712 | -83.58
1.4716 | -83.5751
1.4719 | -83.5697
1.4723 | -83.5648
1.4726 | -83.556
1.473 | -83.5496
1.4733 | -83.5434
1.4736 | -83.533
1.474 | -83.5242
1.4743 | -83.5174
1.4747 | -83.5111
1.475 | -83.506
1.4753 | -83.4995
1.4757 | -83.4932
1.476 | -83.4871
1.4764 | -83.4804
1.4767 | -83.4751
1.477 | -83.4703
1.4774 | -83.4636
1.4777 | -83.4578
1.4781 | -83.4498
1.4784 | -83.4373
1.4787 | -83.4213
1.4791 | -83.4065
1.4794 | -83.3972
1.4798 | -83.3918
1.4801 | -83.3876
1.4804 | -83.3803
1.4808 | -83.3707
1.4811 | -83.3609
1.4815 | -83.3534
1.4818 | -83.3478
1.4821 | -83.3414
1.4825 | -83.3345
1.4828 | -83.3266
1.4832 | -83.3224
1.4835 | -83.3142
1.4838 | -83.3062
1.4842 | -83.3003
1.4845 | -83.2919
1.4849 | -83.2809
1.4852 | -83.2691
1.4855 | -83.2601
1.4859 | -83.2518
1.4862 | -83.2435
1.4866 | -83.236
1.4869 | -83.2282
1.4873 | -83.2152
1.4876 | -83.205
1.4879 | -83.198
1.4883 | -83.1923
1.4886 | -83.1883
1.489 | -83.1825
1.4893 | -83.1724
1.4896 | -83.1597
1.49 | -83.1473
1.4903 | -83.1367
1.4907 | -83.1296
1.491 | -83.1225
1.4913 | -83.1156
1.4917 | -83.1093
1.492 | -83.1021
1.4924 | -83.0946
1.4927 | -83.0876
1.493 | -83.0821
1.4934 | -83.0736
1.4937 | -83.0668
1.4941 | -83.0619
1.4944 | -83.0553
1.4947 | -83.0481
1.4951 | -83.0401
1.4954 | -83.0315
1.4958 | -83.0221
1.4961 | -83.0142
1.4964 | -83.0055
1.4968 | -82.9945
1.4971 | -82.9847
1.4975 | -82.9768
1.4978 | -82.9702
1.4981 | -82.9623
1.4985 | -82.9546
1.4988 | -82.9475
1.4992 | -82.9403
1.4995 | -82.9298
1.4998 | -82.9195
1.5002 | -82.9111
1.5005 | -82.9022
1.5009 | -82.8945
1.5012 | -82.8867
1.5016 | -82.8786
1.5019 | -82.8717
1.5022 | -82.8688
1.5026 | -82.8615
1.5029 | -82.8522
1.5033 | -82.8426
1.5036 | -82.8313
1.5039 | -82.824
1.5043 | -82.8186
1.5046 | -82.8109
1.505 | -82.8017
1.5053 | -82.7943
1.5056 | -82.7856
1.506 | -82.777
1.5063 | -82.7721
1.5067 | -82.7682
1.507 | -82.7606
1.5073 | -82.7522
1.5077 | -82.7457
1.508 | -82.7382
1.5084 | -82.7289
1.5087 | -82.72
1.509 | -82.714
1.5094 | -82.7034
1.5097 | -82.6942
1.5101 | -82.6883
1.5104 | -82.68
1.5107 | -82.6694
1.5111 | -82.6592
1.5114 | -82.6501
1.5118 | -82.6432
1.5121 | -82.635
1.5124 | -82.6263
1.5128 | -82.6192
1.5131 | -82.6114
1.5135 | -82.6046
1.5138 | -82.5972
1.5142 | -82.5904
1.5145 | -82.5813
1.5148 | -82.57
1.5152 | -82.561
1.5155 | -82.5536
1.5159 | -82.5464
1.5162 | -82.5401
1.5165 | -82.5327
1.5169 | -82.524
1.5172 | -82.5155
1.5176 | -82.5073
1.5179 | -82.5002
1.5182 | -82.4936
1.5186 | -82.4872
1.5189 | -82.4808
1.5193 | -82.4737
1.5196 | -82.4668
1.5199 | -82.4608
1.5203 | -82.4552
1.5206 | -82.4489
1.521 | -82.4387
1.5213 | -82.4319
1.5216 | -82.4258
1.522 | -82.4201
1.5223 | -82.4132
1.5227 | -82.4098
1.523 | -82.4053
1.5233 | -82.3994
1.5237 | -82.3951
1.524 | -82.3891
1.5244 | -82.3841
1.5247 | -82.3805
1.525 | -82.3758
1.5254 | -82.3711
1.5257 | -82.3668
1.5261 | -82.3627
1.5264 | -82.3577
1.5267 | -82.3515
1.5271 | -82.3459
1.5274 | -82.3414
1.5278 | -82.3382
1.5281 | -82.336
1.5285 | -82.3373
1.5288 | -82.3363
1.5291 | -82.3322
1.5295 | -82.3258
1.5298 | -82.32
1.5302 | -82.3151
1.5305 | -82.3098
1.5308 | -82.3034
1.5312 | -82.2987
1.5315 | -82.2962
1.5319 | -82.2969
1.5322 | -82.2991
1.5325 | -82.2982
1.5329 | -82.2959
1.5332 | -82.2953
1.5336 | -82.2956
1.5339 | -82.295
1.5342 | -82.2935
1.5346 | -82.2957
1.5349 | -82.2985
1.5353 | -82.2977
1.5356 | -82.2961
1.5359 | -82.2952
1.5363 | -82.2938
1.5366 | -82.2916
1.537 | -82.2899
1.5373 | -82.2903
1.5376 | -82.2911
1.538 | -82.2915
1.5383 | -82.2948
1.5387 | -82.2991
1.539 | -82.3003
1.5393 | -82.3019
1.5397 | -82.3047
1.54 | -82.3059
1.5404 | -82.3043
1.5407 | -82.3026
1.541 | -82.3
1.5414 | -82.2989
1.5417 | -82.2981
1.5421 | -82.3013
1.5424 | -82.3018
1.5428 | -82.3022
1.5431 | -82.3017
1.5434 | -82.3009
1.5438 | -82.3023
1.5441 | -82.3065
1.5445 | -82.309
1.5448 | -82.3113
1.5451 | -82.3134
1.5455 | -82.3164
1.5458 | -82.3199
1.5462 | -82.3239
1.5465 | -82.3284
1.5468 | -82.3319
1.5472 | -82.3344
1.5475 | -82.338
1.5479 | -82.3435
1.5482 | -82.3504
1.5485 | -82.3551
1.5489 | -82.3582
1.5492 | -82.364
1.5496 | -82.3713
1.5499 | -82.3784
1.5502 | -82.3832
1.5506 | -82.3873
1.5509 | -82.3911
1.5513 | -82.3948
1.5516 | -82.3987
1.5519 | -82.4036
1.5523 | -82.4093
1.5526 | -82.4141
1.553 | -82.4171
1.5533 | -82.4229
1.5536 | -82.4285
1.554 | -82.433
1.5543 | -82.4385
1.5547 | -82.4428
1.555 | -82.4465
1.5554 | -82.4517
1.5557 | -82.4562
1.556 | -82.4629
1.5564 | -82.4716
1.5567 | -82.479
1.5571 | -82.4843
1.5574 | -82.4897
1.5577 | -82.494
1.5581 | -82.4979
1.5584 | -82.5063
1.5588 | -82.5157
1.5591 | -82.524
1.5594 | -82.5302
1.5598 | -82.5325
1.5601 | -82.5342
1.5605 | -82.5372
1.5608 | -82.5426
1.5611 | -82.5489
1.5615 | -82.5548
1.5618 | -82.5618
1.5622 | -82.5686
1.5625 | -82.575
1.5628 | -82.5817
1.5632 | -82.5887
1.5635 | -82.5962
1.5639 | -82.6035
1.5642 | -82.6106
1.5645 | -82.6171
1.5649 | -82.6235
1.5652 | -82.6284
1.5656 | -82.6339
1.5659 | -82.6409
1.5662 | -82.6474
1.5666 | -82.6529
1.5669 | -82.6595
1.5673 | -82.6676
1.5676 | -82.6744
1.5679 | -82.6813
1.5683 | -82.6905
1.5686 | -82.7001
1.569 | -82.7063
1.5693 | -82.712
1.5696 | -82.7214
1.57 | -82.7282
1.5703 | -82.7328
1.5707 | -82.7409
1.571 | -82.7476
1.5714 | -82.7536
1.5717 | -82.7506
1.572 | -82.7546
1.5724 | -82.7614
1.5727 | -82.7669
1.5731 | -82.7717
1.5734 | -82.7763
1.5737 | -82.7788
1.5741 | -82.7826
1.5744 | -82.7881
1.5748 | -82.7944
1.5751 | -82.8007
1.5754 | -82.8042
1.5758 | -82.8081
1.5761 | -82.8133
1.5765 | -82.8174
1.5768 | -82.8218
1.5771 | -82.8283
1.5775 | -82.8368
1.5778 | -82.8416
1.5782 | -82.8469
1.5785 | -82.8508
1.5788 | -82.8501
1.5792 | -82.8477
1.5795 | -82.8465
1.5799 | -82.8483
1.5802 | -82.8532
1.5805 | -82.8574
1.5809 | -82.8611
1.5812 | -82.8654
1.5816 | -82.8652
1.5819 | -82.8683
1.5822 | -82.8725
1.5826 | -82.8752
1.5829 | -82.8734
1.5833 | -82.8735
1.5836 | -82.8756
1.584 | -82.8777
1.5843 | -82.8803
1.5846 | -82.882
1.585 | -82.8847
1.5853 | -82.8893
1.5857 | -82.8975
1.586 | -82.9054
1.5863 | -82.9093
1.5867 | -82.9082
1.587 | -82.9072
1.5874 | -82.9092
1.5877 | -82.913
1.588 | -82.9163
1.5884 | -82.9197
1.5887 | -82.921
1.5891 | -82.9213
1.5894 | -82.9227
1.5897 | -82.9237
1.5901 | -82.9254
1.5904 | -82.9252
1.5908 | -82.9275
1.5911 | -82.9296
1.5914 | -82.9309
1.5918 | -82.9321
1.5921 | -82.9358
1.5925 | -82.9411
1.5928 | -82.9467
1.5931 | -82.9509
1.5935 | -82.9519
1.5938 | -82.9522
1.5942 | -82.9534
1.5945 | -82.9549
1.5948 | -82.9618
1.5952 | -82.9639
1.5955 | -82.9645
1.5959 | -82.9628
1.5962 | -82.9588
1.5965 | -82.9579
1.5969 | -82.96
1.5972 | -82.9651
1.5976 | -82.9693
1.5979 | -82.9709
1.5983 | -82.9696
1.5986 | -82.9673
1.5989 | -82.9665
1.5993 | -82.9663
1.5996 | -82.9668
1.6 | -82.968
1.6003 | -82.9685
1.6006 | -82.9685
1.601 | -82.9709
1.6013 | -82.9723
1.6017 | -82.9737
1.602 | -82.973
1.6023 | -82.9727
1.6027 | -82.9745
1.603 | -82.9759
1.6034 | -82.9774
1.6037 | -82.9755
1.604 | -82.9748
1.6044 | -82.9784
1.6047 | -82.984
1.6051 | -82.9875
1.6054 | -82.9888
1.6057 | -82.9881
1.6061 | -82.988
1.6064 | -82.9877
1.6068 | -82.9835
1.6071 | -82.9813
1.6074 | -82.9819
1.6078 | -82.9851
1.6081 | -82.9879
1.6085 | -82.9883
1.6088 | -82.9889
1.6091 | -82.9901
1.6095 | -82.992
1.6098 | -82.996
1.6102 | -82.9971
1.6105 | -82.9977
1.6108 | -82.9993
1.6112 | -82.999
1.6115 | -83
1.6119 | -83.0018
1.6122 | -83.0026
1.6126 | -83.0058
1.6129 | -83.008
1.6132 | -83.0071
1.6136 | -83.0034
1.6139 | -83.0005
1.6143 | -82.9982
1.6146 | -82.9988
1.6149 | -83.0003
1.6153 | -82.9971
1.6156 | -82.9975
1.616 | -82.9989
1.6163 | -82.9992
1.6166 | -82.9997
1.617 | -83.0003
1.6173 | -83.0033
1.6177 | -83.0081
1.618 | -83.01
1.6183 | -83.0133
1.6187 | -83.0172
1.619 | -83.0191
1.6194 | -83.0224
1.6197 | -83.0262
1.62 | -83.0283
1.6204 | -83.0316
1.6207 | -83.0379
1.6211 | -83.0413
1.6214 | -83.0445
1.6217 | -83.0473
1.6221 | -83.0479
1.6224 | -83.0448
1.6228 | -83.045
1.6231 | -83.0445
1.6234 | -83.0454
1.6238 | -83.0454
1.6241 | -83.0482
1.6245 | -83.0543
1.6248 | -83.0582
1.6252 | -83.0551
1.6255 | -83.0523
1.6258 | -83.054
1.6262 | -83.0542
1.6265 | -83.0521
1.6269 | -83.0527
1.6272 | -83.0551
1.6275 | -83.0554
1.6279 | -83.0541
1.6282 | -83.0537
1.6286 | -83.056
1.6289 | -83.0592
1.6292 | -83.0626
1.6296 | -83.0639
1.6299 | -83.0662
1.6303 | -83.0703
1.6306 | -83.0727
1.6309 | -83.0815
1.6313 | -83.0867
1.6316 | -83.0894
1.632 | -83.0917
1.6323 | -83.094
1.6326 | -83.0966
1.633 | -83.0993
1.6333 | -83.1011
1.6337 | -83.1018
1.634 | -83.1023
1.6343 | -83.1037
1.6347 | -83.1052
1.635 | -83.1077
1.6354 | -83.1099
1.6357 | -83.1124
1.636 | -83.114
1.6364 | -83.1135
1.6367 | -83.1168
1.6371 | -83.1205
1.6374 | -83.1223
1.6377 | -83.1234
1.6381 | -83.1221
1.6384 | -83.1202
1.6388 | -83.1202
1.6391 | -83.1208
1.6395 | -83.1222
1.6398 | -83.1255
1.6401 | -83.1271
1.6405 | -83.1231
1.6408 | -83.1247
1.6412 | -83.1282
1.6415 | -83.1334
1.6418 | -83.1351
1.6422 | -83.1343
1.6425 | -83.1349
1.6429 | -83.1356
1.6432 | -83.1372
1.6435 | -83.139
1.6439 | -83.1406
1.6442 | -83.14
1.6446 | -83.1397
1.6449 | -83.1406
1.6452 | -83.144
1.6456 | -83.1466
1.6459 | -83.1481
1.6463 | -83.1493
1.6466 | -83.1503
1.6469 | -83.1513
1.6473 | -83.1528
1.6476 | -83.1536
1.648 | -83.1542
1.6483 | -83.1555
1.6486 | -83.1554
1.649 | -83.1564
1.6493 | -83.1559
1.6497 | -83.1569
1.65 | -83.157
1.6503 | -83.1558
1.6507 | -83.1551
1.651 | -83.1537
1.6514 | -83.1536
1.6517 | -83.1546
1.652 | -83.1565
1.6524 | -83.158
1.6527 | -83.1598
1.6531 | -83.1613
1.6534 | -83.1634
1.6538 | -83.163
1.6541 | -83.1617
1.6544 | -83.1616
1.6548 | -83.1635
1.6551 | -83.1654
1.6555 | -83.165
1.6558 | -83.1629
1.6561 | -83.1623
1.6565 | -83.1613
1.6568 | -83.1607
1.6572 | -83.1591
1.6575 | -83.1583
1.6578 | -83.1568
1.6582 | -83.1555
1.6585 | -83.1543
1.6589 | -83.1547
1.6592 | -83.1539
1.6595 | -83.1516
1.6599 | -83.1504
1.6602 | -83.1503
1.6606 | -83.1508
1.6609 | -83.1509
1.6612 | -83.1509
1.6616 | -83.1505
1.6619 | -83.1496
1.6623 | -83.1481
1.6626 | -83.1466
1.6629 | -83.1467
1.6633 | -83.1447
1.6636 | -83.1395
1.664 | -83.1399
1.6643 | -83.1412
1.6646 | -83.1406
1.665 | -83.1394
1.6653 | -83.1394
1.6657 | -83.1382
1.666 | -83.1376
1.6663 | -83.1375
1.6667 | -83.1359
1.667 | -83.1319
1.6674 | -83.1315
1.6677 | -83.1316
1.6681 | -83.1276
1.6684 | -83.1247
1.6687 | -83.1248
1.6691 | -83.1238
1.6694 | -83.1219
1.6698 | -83.1223
1.6701 | -83.1223
1.6704 | -83.1217
1.6708 | -83.1203
1.6711 | -83.1178
1.6715 | -83.1165
1.6718 | -83.1158
1.6721 | -83.1165
1.6725 | -83.1201
1.6728 | -83.1238
1.6732 | -83.1255
1.6735 | -83.1248
1.6738 | -83.1237
1.6742 | -83.1239
1.6745 | -83.1249
1.6749 | -83.1228
1.6752 | -83.1225
1.6755 | -83.1228
1.6759 | -83.123
1.6762 | -83.1223
1.6766 | -83.1225
1.6769 | -83.1227
1.6772 | -83.1239
1.6776 | -83.1235
1.6779 | -83.1235
1.6783 | -83.1255
1.6786 | -83.1275
1.6789 | -83.1279
1.6793 | -83.1284
1.6796 | -83.1303
1.68 | -83.1304
1.6803 | -83.133
1.6806 | -83.1292
1.681 | -83.1242
1.6813 | -83.1254
1.6817 | -83.1268
1.682 | -83.1286
1.6824 | -83.1329
1.6827 | -83.1362
1.683 | -83.1394
1.6834 | -83.1456
1.6837 | -83.15
1.6841 | -83.1519
1.6844 | -83.1542
1.6847 | -83.1569
1.6851 | -83.1626
1.6854 | -83.1712
1.6858 | -83.1776
1.6861 | -83.1807
1.6864 | -83.1865
1.6868 | -83.1911
1.6871 | -83.1952
1.6875 | -83.2008
1.6878 | -83.205
1.6881 | -83.2085
1.6885 | -83.2096
1.6888 | -83.2123
1.6892 | -83.2171
1.6895 | -83.2263
1.6898 | -83.2299
1.6902 | -83.234
1.6905 | -83.2384
1.6909 | -83.242
1.6912 | -83.2461
1.6915 | -83.2504
1.6919 | -83.2554
1.6922 | -83.2603
1.6926 | -83.2665
1.6929 | -83.2788
1.6932 | -83.2911
1.6936 | -83.3013
1.6939 | -83.3087
1.6943 | -83.3146
1.6946 | -83.32
1.6949 | -83.3251
1.6953 | -83.3308
1.6956 | -83.3393
1.696 | -83.3449
1.6963 | -83.3494
1.6967 | -83.3539
1.697 | -83.3592
1.6973 | -83.3663
1.6977 | -83.3687
1.698 | -83.3731
1.6984 | -83.3797
1.6987 | -83.3856
1.699 | -83.3919
1.6994 | -83.3998
1.6997 | -83.407
1.7001 | -83.4169
1.7004 | -83.4253
1.7007 | -83.4318
1.7011 | -83.4359
1.7014 | -83.4409
1.7018 | -83.4496
1.7021 | -83.4574
1.7024 | -83.4638
1.7028 | -83.4699
1.7031 | -83.4737
1.7035 | -83.4746
1.7038 | -83.4809
1.7041 | -83.487
1.7045 | -83.4953
1.7048 | -83.502
1.7052 | -83.5088
1.7055 | -83.5149
1.7058 | -83.5203
1.7062 | -83.5274
1.7065 | -83.5347
1.7069 | -83.5415
1.7072 | -83.547
1.7075 | -83.5542
1.7079 | -83.5594
1.7082 | -83.566
1.7086 | -83.5745
1.7089 | -83.5827
1.7092 | -83.5906
1.7096 | -83.5982
1.7099 | -83.6028
1.7103 | -83.6068
1.7106 | -83.6156
1.711 | -83.6201
1.7113 | -83.6294
1.7116 | -83.6409
1.712 | -83.6491
1.7123 | -83.6569
1.7127 | -83.6642
1.713 | -83.6711
1.7133 | -83.6788
1.7137 | -83.6869
1.714 | -83.6965
1.7144 | -83.7046
1.7147 | -83.7118
1.715 | -83.7193
1.7154 | -83.7265
1.7157 | -83.7336
1.7161 | -83.7411
1.7164 | -83.7481
1.7167 | -83.7545
1.7171 | -83.7609
1.7174 | -83.7701
1.7178 | -83.7778
1.7181 | -83.7828
1.7184 | -83.7887
1.7188 | -83.7948
1.7191 | -83.7988
1.7195 | -83.804
1.7198 | -83.8109
1.7201 | -83.8171
1.7205 | -83.8241
1.7208 | -83.8317
1.7212 | -83.838
1.7215 | -83.8451
1.7218 | -83.8517
1.7222 | -83.857
1.7225 | -83.8614
1.7229 | -83.8693
1.7232 | -83.8774
1.7236 | -83.8802
1.7239 | -83.8842
1.7242 | -83.8867
1.7246 | -83.8897
1.7249 | -83.8935
1.7253 | -83.8979
1.7256 | -83.9026
1.7259 | -83.9073
1.7263 | -83.9133
1.7266 | -83.9183
1.727 | -83.925
1.7273 | -83.9311
1.7276 | -83.9378
1.728 | -83.944
1.7283 | -83.9477
1.7287 | -83.9533
1.729 | -83.9582
1.7293 | -83.9619
1.7297 | -83.9588
1.73 | -83.9579
1.7304 | -83.9644
1.7307 | -83.9716
1.731 | -83.9758
1.7314 | -83.9804
1.7317 | -83.9841
1.7321 | -83.9908
1.7324 | -83.9945
1.7327 | -83.9974
1.7331 | -84.0009
1.7334 | -84.0045
1.7338 | -84.0074
1.7341 | -84.0103
1.7344 | -84.0097
1.7348 | -84.0075
1.7351 | -84.0092
1.7355 | -84.0118
1.7358 | -84.0144
1.7361 | -84.0176
1.7365 | -84.0204
1.7368 | -84.0252
1.7372 | -84.028
1.7375 | -84.0301
1.7379 | -84.0321
1.7382 | -84.0333
1.7385 | -84.0328
1.7389 | -84.0344
1.7392 | -84.0368
1.7396 | -84.0415
1.7399 | -84.0497
1.7402 | -84.055
1.7406 | -84.0596
1.7409 | -84.0627
1.7413 | -84.0677
1.7416 | -84.0717
1.7419 | -84.0729
1.7423 | -84.0728
1.7426 | -84.0788
1.743 | -84.0847
1.7433 | -84.0854
1.7436 | -84.0894
1.744 | -84.0971
1.7443 | -84.1011
1.7447 | -84.1019
1.745 | -84.1044
1.7453 | -84.1089
1.7457 | -84.1115
1.746 | -84.1111
1.7464 | -84.111
1.7467 | -84.1111
1.747 | -84.1117
1.7474 | -84.1101
1.7477 | -84.1109
1.7481 | -84.1139
1.7484 | -84.1142
1.7487 | -84.1121
1.7491 | -84.1102
1.7494 | -84.1094
1.7498 | -84.1097
1.7501 | -84.1097
1.7504 | -84.1109
1.7508 | -84.1119
1.7511 | -84.1109
1.7515 | -84.1117
1.7518 | -84.1152
1.7522 | -84.1187
1.7525 | -84.1194
1.7528 | -84.1177
1.7532 | -84.1161
1.7535 | -84.1157
1.7539 | -84.1154
1.7542 | -84.1151
1.7545 | -84.118
1.7549 | -84.1183
1.7552 | -84.1194
1.7556 | -84.1209
1.7559 | -84.1207
1.7562 | -84.1229
1.7566 | -84.1271
1.7569 | -84.1278
1.7573 | -84.1285
1.7576 | -84.1307
1.7579 | -84.1314
1.7583 | -84.132
1.7586 | -84.1341
1.759 | -84.136
1.7593 | -84.1346
1.7596 | -84.1311
1.76 | -84.1305
1.7603 | -84.1281
1.7607 | -84.1254
1.761 | -84.1257
1.7613 | -84.1249
1.7617 | -84.1239
1.762 | -84.1216
1.7624 | -84.1206
1.7627 | -84.1167
1.763 | -84.1147
1.7634 | -84.1152
1.7637 | -84.1148
1.7641 | -84.114
1.7644 | -84.1129
1.7648 | -84.1121
1.7651 | -84.1104
1.7654 | -84.1087
1.7658 | -84.1082
1.7661 | -84.1066
1.7665 | -84.1061
1.7668 | -84.1054
1.7671 | -84.1032
1.7675 | -84.1013
1.7678 | -84.0967
1.7682 | -84.0934
1.7685 | -84.0932
1.7688 | -84.0917
1.7692 | -84.0931
1.7695 | -84.0942
1.7699 | -84.0935
1.7702 | -84.0932
1.7705 | -84.092
1.7709 | -84.0918
1.7712 | -84.0931
1.7716 | -84.0899
1.7719 | -84.0909
1.7722 | -84.0908
1.7726 | -84.0893
1.7729 | -84.0886
1.7733 | -84.0896
1.7736 | -84.0897
1.7739 | -84.0889
1.7743 | -84.0873
1.7746 | -84.0856
1.775 | -84.0859
1.7753 | -84.0846
1.7756 | -84.0806
1.776 | -84.0745
1.7763 | -84.0713
1.7767 | -84.0717
1.777 | -84.072
1.7773 | -84.0722
1.7777 | -84.0711
1.778 | -84.0731
1.7784 | -84.0738
1.7787 | -84.074
1.7791 | -84.0745
1.7794 | -84.0759
1.7797 | -84.0759
1.7801 | -84.0771
1.7804 | -84.0774
1.7808 | -84.0778
1.7811 | -84.0792
1.7814 | -84.0803
1.7818 | -84.0816
1.7821 | -84.0804
1.7825 | -84.0762
1.7828 | -84.0746
1.7831 | -84.0746
1.7835 | -84.0774
1.7838 | -84.0833
1.7842 | -84.0854
1.7845 | -84.0852
1.7848 | -84.0856
1.7852 | -84.0879
1.7855 | -84.0901
1.7859 | -84.0915
1.7862 | -84.0914
1.7865 | -84.0898
1.7869 | -84.0889
1.7872 | -84.0916
1.7876 | -84.0944
1.7879 | -84.0984
1.7882 | -84.1005
1.7886 | -84.1063
1.7889 | -84.1092
1.7893 | -84.1095
1.7896 | -84.1113
1.7899 | -84.1134
1.7903 | -84.1153
1.7906 | -84.1164
1.791 | -84.1167
1.7913 | -84.1188
1.7916 | -84.123
1.792 | -84.1266
1.7923 | -84.1287
1.7927 | -84.1299
1.793 | -84.1303
1.7934 | -84.132
1.7937 | -84.1342
1.794 | -84.1373
1.7944 | -84.1401
1.7947 | -84.1414
1.7951 | -84.1421
1.7954 | -84.1424
1.7957 | -84.1416
1.7961 | -84.1411
1.7964 | -84.1427
1.7968 | -84.1442
1.7971 | -84.1472
1.7974 | -84.1513
1.7978 | -84.1539
1.7981 | -84.1578
1.7985 | -84.1592
1.7988 | -84.1602
1.7991 | -84.1626
1.7995 | -84.1636
1.7998 | -84.1619
1.8002 | -84.1648
1.8005 | -84.1671
1.8008 | -84.1709
1.8012 | -84.1745
1.8015 | -84.1792
1.8019 | -84.1783
1.8022 | -84.1833
1.8025 | -84.1894
1.8029 | -84.1904
1.8032 | -84.1919
1.8036 | -84.1934
1.8039 | -84.1945
1.8042 | -84.196
1.8046 | -84.1982
1.8049 | -84.1999
1.8053 | -84.2007
1.8056 | -84.2021
1.8059 | -84.2041
1.8063 | -84.2064
1.8066 | -84.2073
1.807 | -84.2059
1.8073 | -84.2079
1.8077 | -84.2093
1.808 | -84.2107
1.8083 | -84.2095
1.8087 | -84.2083
1.809 | -84.2112
1.8094 | -84.2152
1.8097 | -84.2171
1.81 | -84.216
1.8104 | -84.2163
1.8107 | -84.2173
1.8111 | -84.2187
1.8114 | -84.22
1.8117 | -84.2218
1.8121 | -84.224
1.8124 | -84.2261
1.8128 | -84.227
1.8131 | -84.2301
1.8134 | -84.2314
1.8138 | -84.2317
1.8141 | -84.2322
1.8145 | -84.2326
1.8148 | -84.2357
1.8151 | -84.2376
1.8155 | -84.239
1.8158 | -84.2401
1.8162 | -84.2422
1.8165 | -84.2412
1.8168 | -84.2414
1.8172 | -84.2451
1.8175 | -84.247
1.8179 | -84.2485
1.8182 | -84.2485
1.8185 | -84.2509
1.8189 | -84.2545
1.8192 | -84.2581
1.8196 | -84.2617
1.8199 | -84.2676
1.8202 | -84.2709
1.8206 | -84.2706
1.8209 | -84.2717
1.8213 | -84.2728
1.8216 | -84.2754
1.822 | -84.2778
1.8223 | -84.2814
1.8226 | -84.284
1.823 | -84.2852
1.8233 | -84.2862
1.8237 | -84.2881
1.824 | -84.2908
1.8243 | -84.295
1.8247 | -84.2976
1.825 | -84.2992
1.8254 | -84.3009
1.8257 | -84.3032
1.826 | -84.3079
1.8264 | -84.3155
1.8267 | -84.3181
1.8271 | -84.3217
1.8274 | -84.3225
1.8277 | -84.3238
1.8281 | -84.3254
1.8284 | -84.3281
1.8288 | -84.3299
1.8291 | -84.328
1.8294 | -84.3261
1.8298 | -84.3264
1.8301 | -84.328
1.8305 | -84.3317
1.8308 | -84.332
1.8311 | -84.3325
1.8315 | -84.3334
1.8318 | -84.3338
1.8322 | -84.3343
1.8325 | -84.3345
1.8328 | -84.3373
1.8332 | -84.3384
1.8335 | -84.3423
1.8339 | -84.3436
1.8342 | -84.3399
1.8345 | -84.3419
1.8349 | -84.3436
1.8352 | -84.3455
1.8356 | -84.3476
1.8359 | -84.3482
1.8363 | -84.3491
1.8366 | -84.3504
1.8369 | -84.3527
1.8373 | -84.3536
1.8376 | -84.353
1.838 | -84.3511
1.8383 | -84.3508
1.8386 | -84.3507
1.839 | -84.3508
1.8393 | -84.3521
1.8397 | -84.3531
1.84 | -84.3535
1.8403 | -84.3556
1.8407 | -84.356
1.841 | -84.3577
1.8414 | -84.3573
1.8417 | -84.3608
1.842 | -84.3643
1.8424 | -84.3649
1.8427 | -84.3666
1.8431 | -84.3691
1.8434 | -84.372
1.8437 | -84.3738
1.8441 | -84.3787
1.8444 | -84.3848
1.8448 | -84.3851
1.8451 | -84.385
1.8454 | -84.3887
1.8458 | -84.389
1.8461 | -84.3893
1.8465 | -84.392
1.8468 | -84.3912
1.8471 | -84.3953
1.8475 | -84.3993
1.8478 | -84.4005
1.8482 | -84.4027
1.8485 | -84.4043
1.8489 | -84.4073
1.8492 | -84.4078
1.8495 | -84.4038
1.8499 | -84.4029
1.8502 | -84.4032
1.8506 | -84.4033
1.8509 | -84.4035
1.8512 | -84.404
1.8516 | -84.4053
1.8519 | -84.4059
1.8523 | -84.4045
1.8526 | -84.4033
1.8529 | -84.4054
1.8533 | -84.4055
1.8536 | -84.4031
1.854 | -84.3999
1.8543 | -84.4012
1.8546 | -84.4096
1.855 | -84.4172
1.8553 | -84.4202
1.8557 | -84.4228
1.856 | -84.4255
1.8563 | -84.4269
1.8567 | -84.4278
1.857 | -84.4291
1.8574 | -84.4297
1.8577 | -84.4285
1.858 | -84.4287
1.8584 | -84.4293
1.8587 | -84.4251
1.8591 | -84.4236
1.8594 | -84.4235
1.8597 | -84.4264
1.8601 | -84.4285
1.8604 | -84.4283
1.8608 | -84.428
1.8611 | -84.4252
1.8614 | -84.4238
1.8618 | -84.4235
1.8621 | -84.4194
1.8625 | -84.4186
1.8628 | -84.4184
1.8632 | -84.4177
1.8635 | -84.4194
1.8638 | -84.4198
1.8642 | -84.4208
1.8645 | -84.4194
1.8649 | -84.4174
1.8652 | -84.4176
1.8655 | -84.417
1.8659 | -84.4159
1.8662 | -84.4156
1.8666 | -84.4148
1.8669 | -84.4132
1.8672 | -84.411
1.8676 | -84.4103
1.8679 | -84.408
1.8683 | -84.4078
1.8686 | -84.4045
1.8689 | -84.399
1.8693 | -84.3984
1.8696 | -84.398
1.87 | -84.3972
1.8703 | -84.3966
1.8706 | -84.3953
1.871 | -84.3947
1.8713 | -84.3941
1.8717 | -84.3941
1.872 | -84.3937
1.8723 | -84.3896
1.8727 | -84.3878
1.873 | -84.3877
1.8734 | -84.3889
1.8737 | -84.3883
1.874 | -84.3871
1.8744 | -84.386
1.8747 | -84.3854
1.8751 | -84.3853
1.8754 | -84.3848
1.8757 | -84.3834
1.8761 | -84.3842
1.8764 | -84.3829
1.8768 | -84.38
1.8771 | -84.3761
1.8775 | -84.3734
1.8778 | -84.3727
1.8781 | -84.3695
1.8785 | -84.3648
1.8788 | -84.3576
1.8792 | -84.3553
1.8795 | -84.3556
1.8798 | -84.3547
1.8802 | -84.3522
1.8805 | -84.3469
1.8809 | -84.3443
1.8812 | -84.3412
1.8815 | -84.337
1.8819 | -84.3335
1.8822 | -84.3197
1.8826 | -84.3073
1.8829 | -84.3035
1.8832 | -84.3008
1.8836 | -84.2982
1.8839 | -84.2924
1.8843 | -84.286
1.8846 | -84.2807
1.8849 | -84.2783
1.8853 | -84.2755
1.8856 | -84.271
1.886 | -84.2638
1.8863 | -84.2627
1.8866 | -84.2623
1.887 | -84.2617
1.8873 | -84.2573
1.8877 | -84.2548
1.888 | -84.2547
1.8883 | -84.2518
1.8887 | -84.2478
1.889 | -84.2441
1.8894 | -84.2415
1.8897 | -84.2376
1.8901 | -84.2368
1.8904 | -84.2357
1.8907 | -84.2304
1.8911 | -84.2258
1.8914 | -84.2221
1.8918 | -84.2161
1.8921 | -84.2098
1.8924 | -84.2054
1.8928 | -84.2012
1.8931 | -84.1957
1.8935 | -84.1922
1.8938 | -84.1889
1.8941 | -84.1855
1.8945 | -84.1813
1.8948 | -84.1781
1.8952 | -84.1763
1.8955 | -84.1761
1.8958 | -84.1686
1.8962 | -84.1623
1.8965 | -84.1549
1.8969 | -84.1485
1.8972 | -84.1414
1.8975 | -84.1331
1.8979 | -84.1238
1.8982 | -84.1154
1.8986 | -84.1102
1.8989 | -84.1065
1.8992 | -84.1009
1.8996 | -84.0983
1.8999 | -84.0932
1.9003 | -84.0871
1.9006 | -84.0828
1.9009 | -84.0774
1.9013 | -84.072
1.9016 | -84.0671
1.902 | -84.0626
1.9023 | -84.0581
1.9026 | -84.052
1.903 | -84.0455
1.9033 | -84.0414
1.9037 | -84.0381
1.904 | -84.0265
1.9043 | -84.0139
1.9047 | -84.0056
1.905 | -83.9984
1.9054 | -83.9908
1.9057 | -83.9832
1.9061 | -83.9771
1.9064 | -83.9712
1.9067 | -83.9647
1.9071 | -83.9583
1.9074 | -83.9512
1.9078 | -83.9427
1.9081 | -83.9372
1.9084 | -83.9324
1.9088 | -83.9258
1.9091 | -83.9189
1.9095 | -83.9125
1.9098 | -83.9059
1.9101 | -83.8983
1.9105 | -83.8908
1.9108 | -83.886
1.9112 | -83.8816
1.9115 | -83.8763
1.9118 | -83.8696
1.9122 | -83.8616
1.9125 | -83.8543
1.9129 | -83.8473
1.9132 | -83.8397
1.9135 | -83.8266
1.9139 | -83.817
1.9142 | -83.8089
1.9146 | -83.7973
1.9149 | -83.7894
1.9152 | -83.7847
1.9156 | -83.7784
1.9159 | -83.7726
1.9163 | -83.7654
1.9166 | -83.7582
1.9169 | -83.751
1.9173 | -83.7436
1.9176 | -83.7368
1.918 | -83.7284
1.9183 | -83.7211
1.9187 | -83.7143
1.919 | -83.7064
1.9193 | -83.6999
1.9197 | -83.6879
1.92 | -83.6768
1.9204 | -83.6699
1.9207 | -83.662
1.921 | -83.6533
1.9214 | -83.6427
1.9217 | -83.6327
1.9221 | -83.6238
1.9224 | -83.6164
1.9227 | -83.6094
1.9231 | -83.6015
1.9234 | -83.5941
1.9238 | -83.5874
1.9241 | -83.5813
1.9244 | -83.5753
1.9248 | -83.5691
1.9251 | -83.5661
1.9255 | -83.5581
1.9258 | -83.5526
1.9261 | -83.5454
1.9265 | -83.5386
1.9268 | -83.5316
1.9272 | -83.5244
1.9275 | -83.5167
1.9278 | -83.5086
1.9282 | -83.501
1.9285 | -83.4946
1.9289 | -83.4902
1.9292 | -83.4816
1.9295 | -83.4713
1.9299 | -83.4634
1.9302 | -83.4555
1.9306 | -83.45
1.9309 | -83.4418
1.9312 | -83.4344
1.9316 | -83.4285
1.9319 | -83.4231
1.9323 | -83.4181
1.9326 | -83.4133
1.933 | -83.4072
1.9333 | -83.3998
1.9336 | -83.3928
1.934 | -83.3865
1.9343 | -83.3799
1.9347 | -83.3736
1.935 | -83.3673
1.9353 | -83.3549
1.9357 | -83.3445
1.936 | -83.3376
1.9364 | -83.331
1.9367 | -83.324
1.937 | -83.3176
1.9374 | -83.3134
1.9377 | -83.3079
1.9381 | -83.3029
1.9384 | -83.2977
1.9387 | -83.2922
1.9391 | -83.2893
1.9394 | -83.286
1.9398 | -83.2751
1.9401 | -83.2654
1.9404 | -83.2584
1.9408 | -83.2543
1.9411 | -83.2516
1.9415 | -83.2471
1.9418 | -83.24
1.9421 | -83.2304
1.9425 | -83.2228
1.9428 | -83.2165
1.9432 | -83.2063
1.9435 | -83.1968
1.9438 | -83.1889
1.9442 | -83.1846
1.9445 | -83.1802
1.9449 | -83.1765
1.9452 | -83.1686
1.9455 | -83.1639
1.9459 | -83.1579
1.9462 | -83.1507
1.9466 | -83.1427
1.9469 | -83.1337
1.9473 | -83.1261
1.9476 | -83.1219
1.9479 | -83.117
1.9483 | -83.1122
1.9486 | -83.1065
1.949 | -83.1005
1.9493 | -83.0956
1.9496 | -83.0918
1.95 | -83.0871
1.9503 | -83.082
1.9507 | -83.0773
1.951 | -83.0725
1.9513 | -83.0684
1.9517 | -83.0682
1.952 | -83.067
1.9524 | -83.0625
1.9527 | -83.0552
1.953 | -83.0474
1.9534 | -83.0418
1.9537 | -83.0353
1.9541 | -83.0295
1.9544 | -83.026
1.9547 | -83.0244
1.9551 | -83.0191
1.9554 | -83.0121
1.9558 | -83.0093
1.9561 | -83.0065
1.9564 | -83.0028
1.9568 | -82.9992
1.9571 | -82.9956
1.9575 | -82.9909
1.9578 | -82.9858
1.9581 | -82.9814
1.9585 | -82.9746
1.9588 | -82.9672
1.9592 | -82.9628
1.9595 | -82.9603
1.9599 | -82.9574
1.9602 | -82.954
1.9605 | -82.9489
1.9609 | -82.9429
1.9612 | -82.9434
1.9616 | -82.943
1.9619 | -82.9409
1.9622 | -82.9383
1.9626 | -82.9346
1.9629 | -82.9312
1.9633 | -82.9288
1.9636 | -82.9239
1.9639 | -82.9216
1.9643 | -82.9201
1.9646 | -82.9183
1.965 | -82.9103
1.9653 | -82.9037
1.9656 | -82.899
1.966 | -82.8957
1.9663 | -82.8947
1.9667 | -82.895
1.967 | -82.8938
1.9673 | -82.8904
1.9677 | -82.8868
1.968 | -82.8825
1.9684 | -82.879
1.9687 | -82.8726
1.969 | -82.8659
1.9694 | -82.861
1.9697 | -82.8556
1.9701 | -82.8537
1.9704 | -82.85
1.9707 | -82.8469
1.9711 | -82.8441
1.9714 | -82.8416
1.9718 | -82.8383
1.9721 | -82.8339
1.9724 | -82.831
1.9728 | -82.8282
1.9731 | -82.8243
1.9735 | -82.8224
1.9738 | -82.8195
1.9742 | -82.8148
1.9745 | -82.8107
1.9748 | -82.8071
1.9752 | -82.8035
1.9755 | -82.8012
1.9759 | -82.7985
1.9762 | -82.7951
1.9765 | -82.7925
1.9769 | -82.7906
1.9772 | -82.788
1.9776 | -82.7839
1.9779 | -82.7797
1.9782 | -82.7774
1.9786 | -82.7747
1.9789 | -82.7713
1.9793 | -82.7684
1.9796 | -82.7649
1.9799 | -82.763
1.9803 | -82.7607
1.9806 | -82.7578
1.981 | -82.7537
1.9813 | -82.7484
1.9816 | -82.7447
1.982 | -82.7406
1.9823 | -82.7373
1.9827 | -82.7356
1.983 | -82.7359
1.9833 | -82.7345
1.9837 | -82.7307
1.984 | -82.7274
1.9844 | -82.7235
1.9847 | -82.7191
1.985 | -82.7159
1.9854 | -82.713
1.9857 | -82.7097
1.9861 | -82.7078
1.9864 | -82.7056
1.9867 | -82.7031
1.9871 | -82.7008
1.9874 | -82.6998
1.9878 | -82.6993
1.9881 | -82.6973
1.9885 | -82.6942
1.9888 | -82.691
1.9891 | -82.6874
1.9895 | -82.684
1.9898 | -82.6797
1.9902 | -82.674
1.9905 | -82.671
1.9908 | -82.6692
1.9912 | -82.6664
1.9915 | -82.6619
1.9919 | -82.6582
1.9922 | -82.6551
1.9925 | -82.6526
1.9929 | -82.6524
1.9932 | -82.65
1.9936 | -82.6466
1.9939 | -82.6438
1.9942 | -82.6425
1.9946 | -82.6402
1.9949 | -82.6335
1.9953 | -82.6264
1.9956 | -82.6228
1.9959 | -82.6206
1.9963 | -82.6185
1.9966 | -82.618
1.997 | -82.6164
1.9973 | -82.6133
1.9976 | -82.6091
1.998 | -82.6065
1.9983 | -82.6043
1.9987 | -82.6025
1.999 | -82.6018
1.9993 | -82.5991
1.9997 | -82.5927
2 | -82.5809
2.0004 | -82.571
2.0007 | -82.5682
2.001 | -82.5665
2.0014 | -82.5644
2.0017 | -82.5623
2.0021 | -82.5603
2.0024 | -82.5566
2.0028 | -82.5548
2.0031 | -82.5528
2.0034 | -82.553
2.0038 | -82.5528
2.0041 | -82.5502
2.0045 | -82.5464
2.0048 | -82.5441
2.0051 | -82.5414
2.0055 | -82.5369
2.0058 | -82.5321
2.0062 | -82.5269
2.0065 | -82.5242
2.0068 | -82.522
2.0072 | -82.519
2.0075 | -82.5179
2.0079 | -82.5154
2.0082 | -82.5136
2.0085 | -82.5095
2.0089 | -82.5063
2.0092 | -82.5031
2.0096 | -82.4987
2.0099 | -82.4961
2.0102 | -82.4951
2.0106 | -82.4941
2.0109 | -82.4926
2.0113 | -82.4905
2.0116 | -82.4888
2.0119 | -82.4862
2.0123 | -82.4858
2.0126 | -82.4866
2.013 | -82.4864
2.0133 | -82.4853
2.0136 | -82.4849
2.014 | -82.485
2.0143 | -82.4856
2.0147 | -82.4867
2.015 | -82.4869
2.0153 | -82.4857
2.0157 | -82.4844
2.016 | -82.4835
2.0164 | -82.4839
2.0167 | -82.4836
2.0171 | -82.4832
2.0174 | -82.4828
2.0177 | -82.4826
2.0181 | -82.4893
2.0184 | -82.4948
2.0188 | -82.4943
2.0191 | -82.4922
2.0194 | -82.4913
2.0198 | -82.4915
2.0201 | -82.4924
2.0205 | -82.4933
2.0208 | -82.4921
2.0211 | -82.4904
2.0215 | -82.491
2.0218 | -82.4914
2.0222 | -82.4898
2.0225 | -82.4892
2.0228 | -82.4874
2.0232 | -82.4809
2.0235 | -82.4775
2.0239 | -82.4763
2.0242 | -82.4761
2.0245 | -82.4752
2.0249 | -82.4739
2.0252 | -82.4696
2.0256 | -82.4676
2.0259 | -82.4672
2.0262 | -82.467
2.0266 | -82.4665
2.0269 | -82.465
2.0273 | -82.4642
2.0276 | -82.4631
2.0279 | -82.4612
2.0283 | -82.4619
2.0286 | -82.4634
2.029 | -82.4642
2.0293 | -82.4648
2.0297 | -82.4656
2.03 | -82.4662
2.0303 | -82.4669
2.0307 | -82.467
2.031 | -82.4659
2.0314 | -82.4657
2.0317 | -82.4666
2.032 | -82.4661
2.0324 | -82.4659
2.0327 | -82.4678
2.0331 | -82.4682
2.0334 | -82.466
2.0337 | -82.4667
2.0341 | -82.4682
2.0344 | -82.4694
2.0348 | -82.4701
2.0351 | -82.4699
2.0354 | -82.4687
2.0358 | -82.4667
2.0361 | -82.4654
2.0365 | -82.4654
2.0368 | -82.464
2.0371 | -82.4639
2.0375 | -82.4633
2.0378 | -82.4632
2.0382 | -82.4603
2.0385 | -82.4579
2.0388 | -82.4583
2.0392 | -82.4562
2.0395 | -82.4555
2.0399 | -82.4568
2.0402 | -82.4574
2.0405 | -82.4572
2.0409 | -82.4556
2.0412 | -82.461
2.0416 | -82.463
2.0419 | -82.4587
2.0422 | -82.4554
2.0426 | -82.4557
2.0429 | -82.4529
2.0433 | -82.4514
2.0436 | -82.4514
2.044 | -82.4478
2.0443 | -82.4415
2.0446 | -82.4354
2.045 | -82.4289
2.0453 | -82.4229
2.0457 | -82.4201
2.046 | -82.4186
2.0463 | -82.4168
2.0467 | -82.4157
2.047 | -82.4148
2.0474 | -82.4122
2.0477 | -82.4093
2.048 | -82.4088
2.0484 | -82.407
2.0487 | -82.4019
2.0491 | -82.399
2.0494 | -82.3975
2.0497 | -82.3959
2.0501 | -82.3927
2.0504 | -82.3907
2.0508 | -82.389
2.0511 | -82.3878
2.0514 | -82.3903
2.0518 | -82.3895
2.0521 | -82.3865
2.0525 | -82.3871
2.0528 | -82.3861
2.0531 | -82.379
2.0535 | -82.3726
2.0538 | -82.3689
2.0542 | -82.3669
2.0545 | -82.3653
2.0548 | -82.3645
2.0552 | -82.3621
2.0555 | -82.3582
2.0559 | -82.3553
2.0562 | -82.3503
2.0565 | -82.3474
2.0569 | -82.345
2.0572 | -82.3415
2.0576 | -82.3385
2.0579 | -82.3362
2.0583 | -82.336
2.0586 | -82.3323
2.0589 | -82.3271
2.0593 | -82.3248
2.0596 | -82.3215
2.06 | -82.3184
2.0603 | -82.3128
2.0606 | -82.3056
2.061 | -82.3021
2.0613 | -82.2988
2.0617 | -82.2948
2.062 | -82.2907
2.0623 | -82.286
2.0627 | -82.2797
2.063 | -82.275
2.0634 | -82.2715
2.0637 | -82.2673
2.064 | -82.2618
2.0644 | -82.2577
2.0647 | -82.2554
2.0651 | -82.252
2.0654 | -82.2464
2.0657 | -82.2394
2.0661 | -82.2334
2.0664 | -82.2277
2.0668 | -82.2234
2.0671 | -82.2194
2.0674 | -82.2151
2.0678 | -82.2107
2.0681 | -82.2052
2.0685 | -82.2017
2.0688 | -82.2009
2.0691 | -82.1979
2.0695 | -82.1916
2.0698 | -82.1843
2.0702 | -82.1789
2.0705 | -82.1735
2.0708 | -82.1687
2.0712 | -82.1638
2.0715 | -82.159
2.0719 | -82.1544
2.0722 | -82.1486
2.0726 | -82.1406
2.0729 | -82.1362
2.0732 | -82.1317
2.0736 | -82.1262
2.0739 | -82.1191
2.0743 | -82.1113
2.0746 | -82.1062
2.0749 | -82.103
2.0753 | -82.1021
2.0756 | -82.0996
2.076 | -82.0932
2.0763 | -82.0854
2.0766 | -82.0809
2.077 | -82.0764
2.0773 | -82.0713
2.0777 | -82.0662
2.078 | -82.0606
2.0783 | -82.0545
2.0787 | -82.0484
2.079 | -82.0429
2.0794 | -82.0368
2.0797 | -82.0322
2.08 | -82.0265
2.0804 | -82.0213
2.0807 | -82.0159
2.0811 | -82.0101
2.0814 | -82.0027
2.0817 | -81.9967
2.0821 | -81.9936
2.0824 | -81.9885
2.0828 | -81.9828
2.0831 | -81.9754
2.0834 | -81.9707
2.0838 | -81.9658
2.0841 | -81.9599
2.0845 | -81.9538
2.0848 | -81.948
2.0851 | -81.9421
2.0855 | -81.9377
2.0858 | -81.9344
2.0862 | -81.9288
2.0865 | -81.9237
2.0869 | -81.9193
2.0872 | -81.916
2.0875 | -81.9128
2.0879 | -81.9084
2.0882 | -81.9037
2.0886 | -81.8984
2.0889 | -81.8911
2.0892 | -81.8842
2.0896 | -81.8767
2.0899 | -81.8684
2.0903 | -81.8612
2.0906 | -81.8555
2.0909 | -81.8456
2.0913 | -81.837
2.0916 | -81.8313
2.092 | -81.8255
2.0923 | -81.82
2.0926 | -81.8142
2.093 | -81.8071
2.0933 | -81.7997
2.0937 | -81.7947
2.094 | -81.7913
2.0943 | -81.7858
2.0947 | -81.7808
2.095 | -81.7765
2.0954 | -81.7695
2.0957 | -81.7625
2.096 | -81.7565
2.0964 | -81.7523
2.0967 | -81.7491
2.0971 | -81.7438
2.0974 | -81.7387
2.0977 | -81.7325
2.0981 | -81.7264
2.0984 | -81.721
2.0988 | -81.7168
2.0991 | -81.7119
2.0995 | -81.7063
2.0998 | -81.7014
2.1001 | -81.6981
2.1005 | -81.694
2.1008 | -81.6892
2.1012 | -81.6843
2.1015 | -81.68
2.1018 | -81.6769
2.1022 | -81.6749
2.1025 | -81.6726
2.1029 | -81.6682
2.1032 | -81.6635
2.1035 | -81.6592
2.1039 | -81.6549
2.1042 | -81.6509
2.1046 | -81.6476
2.1049 | -81.645
2.1052 | -81.6433
2.1056 | -81.6398
2.1059 | -81.6364
2.1063 | -81.6328
2.1066 | -81.6303
2.1069 | -81.6312
2.1073 | -81.6302
2.1076 | -81.6264
2.108 | -81.6222
2.1083 | -81.618
2.1086 | -81.6134
2.109 | -81.6103
2.1093 | -81.608
2.1097 | -81.6054
2.11 | -81.6037
2.1103 | -81.6013
2.1107 | -81.597
2.111 | -81.5965
2.1114 | -81.596
2.1117 | -81.5951
2.112 | -81.5969
2.1124 | -81.5946
2.1127 | -81.5936
2.1131 | -81.5928
2.1134 | -81.5916
2.1138 | -81.5902
2.1141 | -81.5869
2.1144 | -81.5828
2.1148 | -81.5793
2.1151 | -81.5776
2.1155 | -81.5769
2.1158 | -81.5761
2.1161 | -81.5748
2.1165 | -81.5745
2.1168 | -81.5732
2.1172 | -81.5726
2.1175 | -81.5726
2.1178 | -81.5729
2.1182 | -81.5724
2.1185 | -81.5718
2.1189 | -81.5722
2.1192 | -81.5722
2.1195 | -81.5712
2.1199 | -81.571
2.1202 | -81.5713
2.1206 | -81.5726
2.1209 | -81.5736
2.1212 | -81.5743
2.1216 | -81.5755
2.1219 | -81.5816
2.1223 | -81.5815
2.1226 | -81.581
2.1229 | -81.5822
2.1233 | -81.5835
2.1236 | -81.5852
2.124 | -81.5861
2.1243 | -81.5875
2.1246 | -81.5901
2.125 | -81.5932
2.1253 | -81.5958
2.1257 | -81.5984
2.126 | -81.6011
2.1263 | -81.6023
2.1267 | -81.6048
2.127 | -81.6082
2.1274 | -81.6117
2.1277 | -81.6149
2.1281 | -81.6175
2.1284 | -81.6191
2.1287 | -81.6198
2.1291 | -81.6225
2.1294 | -81.6255
2.1298 | -81.6277
2.1301 | -81.6313
2.1304 | -81.6355
2.1308 | -81.6393
2.1311 | -81.6418
2.1315 | -81.6464
2.1318 | -81.6506
2.1321 | -81.655
2.1325 | -81.6593
2.1328 | -81.664
2.1332 | -81.6671
2.1335 | -81.6717
2.1338 | -81.6782
2.1342 | -81.6821
2.1345 | -81.6842
2.1349 | -81.6886
2.1352 | -81.6954
2.1355 | -81.7016
2.1359 | -81.7071
2.1362 | -81.712
2.1366 | -81.716
2.1369 | -81.7218
2.1372 | -81.7279
2.1376 | -81.7331
2.1379 | -81.738
2.1383 | -81.7429
2.1386 | -81.7491
2.1389 | -81.7582
2.1393 | -81.7664
2.1396 | -81.7722
2.14 | -81.7787
2.1403 | -81.7839
2.1406 | -81.7894
2.141 | -81.7943
2.1413 | -81.8023
2.1417 | -81.8086
2.142 | -81.8167
2.1424 | -81.8231
2.1427 | -81.8285
2.143 | -81.8329
2.1434 | -81.8392
2.1437 | -81.8471
2.1441 | -81.8567
2.1444 | -81.8707
2.1447 | -81.883
2.1451 | -81.8882
2.1454 | -81.8936
2.1458 | -81.9
2.1461 | -81.9071
2.1464 | -81.9146
2.1468 | -81.9216
2.1471 | -81.9272
2.1475 | -81.9348
2.1478 | -81.9418
2.1481 | -81.9477
2.1485 | -81.9536
2.1488 | -81.9624
2.1492 | -81.9699
2.1495 | -81.9768
2.1498 | -81.9847
2.1502 | -81.9937
2.1505 | -82.0028
2.1509 | -82.0108
2.1512 | -82.019
2.1515 | -82.0254
2.1519 | -82.0303
2.1522 | -82.0364
2.1526 | -82.043
2.1529 | -82.0501
2.1532 | -82.0568
2.1536 | -82.0661
2.1539 | -82.0747
2.1543 | -82.0825
2.1546 | -82.0896
2.1549 | -82.091
2.1553 | -82.0928
2.1556 | -82.0974
2.156 | -82.1038
2.1563 | -82.1109
2.1567 | -82.1168
2.157 | -82.1237
2.1573 | -82.1309
2.1577 | -82.1382
2.158 | -82.1453
2.1584 | -82.1528
2.1587 | -82.1599
2.159 | -82.1678
2.1594 | -82.1717
2.1597 | -82.1758
2.1601 | -82.1833
2.1604 | -82.1896
2.1607 | -82.1964
2.1611 | -82.2035
2.1614 | -82.2104
2.1618 | -82.2175
2.1621 | -82.223
2.1624 | -82.2269
2.1628 | -82.2315
2.1631 | -82.238
2.1635 | -82.2437
2.1638 | -82.2498
2.1641 | -82.2565
2.1645 | -82.2615
2.1648 | -82.2678
2.1652 | -82.2734
2.1655 | -82.2801
2.1658 | -82.2863
2.1662 | -82.292
2.1665 | -82.2988
2.1669 | -82.3049
2.1672 | -82.3101
2.1675 | -82.3171
2.1679 | -82.3219
2.1682 | -82.3275
2.1686 | -82.3343
2.1689 | -82.3461
2.1692 | -82.3553
2.1696 | -82.3617
2.1699 | -82.3664
2.1703 | -82.371
2.1706 | -82.3757
2.171 | -82.3799
2.1713 | -82.3861
2.1716 | -82.3945
2.172 | -82.3986
2.1723 | -82.4004
2.1727 | -82.4047
2.173 | -82.4078
2.1733 | -82.4114
2.1737 | -82.4172
2.174 | -82.4221
2.1744 | -82.4268
2.1747 | -82.4298
2.175 | -82.4323
2.1754 | -82.4356
2.1757 | -82.438
2.1761 | -82.4402
2.1764 | -82.4408
2.1767 | -82.4427
2.1771 | -82.4474
2.1774 | -82.4533
2.1778 | -82.4547
2.1781 | -82.4542
2.1784 | -82.458
2.1788 | -82.4618
2.1791 | -82.4634
2.1795 | -82.4664
2.1798 | -82.4713
2.1801 | -82.4755
2.1805 | -82.4791
2.1808 | -82.4773
2.1812 | -82.4784
2.1815 | -82.4796
2.1818 | -82.4825
2.1822 | -82.4873
2.1825 | -82.4914
2.1829 | -82.4957
2.1832 | -82.4975
2.1835 | -82.4973
2.1839 | -82.4994
2.1842 | -82.5004
2.1846 | -82.5027
2.1849 | -82.5037
2.1853 | -82.5076
2.1856 | -82.511
2.1859 | -82.5119
2.1863 | -82.5147
2.1866 | -82.5166
2.187 | -82.5166
2.1873 | -82.5163
2.1876 | -82.5164
2.188 | -82.5165
2.1883 | -82.5166
2.1887 | -82.5206
2.189 | -82.5255
2.1893 | -82.5291
2.1897 | -82.5285
2.19 | -82.5305
2.1904 | -82.5321
2.1907 | -82.5253
2.191 | -82.5231
2.1914 | -82.5229
2.1917 | -82.5257
2.1921 | -82.5297
2.1924 | -82.5335
2.1927 | -82.5378
2.1931 | -82.5392
2.1934 | -82.5383
2.1938 | -82.5344
2.1941 | -82.5301
2.1944 | -82.5283
2.1948 | -82.5261
2.1951 | -82.5257
2.1955 | -82.5252
2.1958 | -82.5227
2.1961 | -82.5222
2.1965 | -82.5239
2.1968 | -82.5238
2.1972 | -82.5228
2.1975 | -82.5212
2.1979 | -82.5176
2.1982 | -82.513
2.1985 | -82.5106
2.1989 | -82.5086
2.1992 | -82.5066
2.1996 | -82.5052
2.1999 | -82.5041
2.2002 | -82.4978
2.2006 | -82.495
2.2009 | -82.4906
2.2013 | -82.4877
2.2016 | -82.4888
2.2019 | -82.4895
2.2023 | -82.4873
2.2026 | -82.4842
2.203 | -82.48
2.2033 | -82.4776
2.2036 | -82.4748
2.204 | -82.4732
2.2043 | -82.4708
2.2047 | -82.4693
2.205 | -82.4675
2.2053 | -82.465
2.2057 | -82.4627
2.206 | -82.4609
2.2064 | -82.4583
2.2067 | -82.4562
2.207 | -82.4568
2.2074 | -82.4559
2.2077 | -82.4525
2.2081 | -82.4502
2.2084 | -82.4492
2.2087 | -82.4501
2.2091 | -82.4502
2.2094 | -82.447
2.2098 | -82.4435
2.2101 | -82.4407
2.2104 | -82.4373
2.2108 | -82.4347
2.2111 | -82.4321
2.2115 | -82.4306
2.2118 | -82.4277
2.2122 | -82.4235
2.2125 | -82.4189
2.2128 | -82.4155
2.2132 | -82.4133
2.2135 | -82.4099
2.2139 | -82.405
2.2142 | -82.4013
2.2145 | -82.3999
2.2149 | -82.3994
2.2152 | -82.3985
2.2156 | -82.3961
2.2159 | -82.3932
2.2162 | -82.3901
2.2166 | -82.3834
2.2169 | -82.3782
2.2173 | -82.3741
2.2176 | -82.3713
2.2179 | -82.3683
2.2183 | -82.3672
2.2186 | -82.3663
2.219 | -82.3643
2.2193 | -82.3632
2.2196 | -82.3622
2.22 | -82.3622
2.2203 | -82.3595
2.2207 | -82.3577
2.221 | -82.3549
2.2213 | -82.3493
2.2217 | -82.3453
2.222 | -82.343
2.2224 | -82.341
2.2227 | -82.3366
2.223 | -82.3331
2.2234 | -82.3311
2.2237 | -82.3303
2.2241 | -82.3303
2.2244 | -82.3284
2.2247 | -82.3261
2.2251 | -82.3233
2.2254 | -82.3216
2.2258 | -82.3213
2.2261 | -82.3209
2.2265 | -82.3173
2.2268 | -82.315
2.2271 | -82.3114
2.2275 | -82.3089
2.2278 | -82.3061
2.2282 | -82.3043
2.2285 | -82.3019
2.2288 | -82.2996
2.2292 | -82.2986
2.2295 | -82.2997
2.2299 | -82.3013
2.2302 | -82.3012
2.2305 | -82.2999
2.2309 | -82.2982
2.2312 | -82.2967
2.2316 | -82.2966
2.2319 | -82.2959
2.2322 | -82.295
2.2326 | -82.2924
2.2329 | -82.2903
2.2333 | -82.2904
2.2336 | -82.2894
2.2339 | -82.2883
2.2343 | -82.2891
2.2346 | -82.2894
2.235 | -82.2881
2.2353 | -82.287
2.2356 | -82.2859
2.236 | -82.2853
2.2363 | -82.2839
2.2367 | -82.2834
2.237 | -82.2823
2.2373 | -82.2817
2.2377 | -82.2813
2.238 | -82.2809
2.2384 | -82.2791
2.2387 | -82.2775
2.2391 | -82.2767
2.2394 | -82.2766
2.2397 | -82.2786
2.2401 | -82.2797
2.2404 | -82.2816
2.2408 | -82.2842
2.2411 | -82.286
2.2414 | -82.2857
2.2418 | -82.2904
2.2421 | -82.2938
2.2425 | -82.2937
2.2428 | -82.2944
2.2431 | -82.2959
2.2435 | -82.2966
2.2438 | -82.2975
2.2442 | -82.2996
2.2445 | -82.3029
2.2448 | -82.3056
2.2452 | -82.3067
2.2455 | -82.3074
2.2459 | -82.3084
2.2462 | -82.3095
2.2465 | -82.3118
2.2469 | -82.315
2.2472 | -82.3168
2.2476 | -82.3169
2.2479 | -82.3166
2.2482 | -82.3163
2.2486 | -82.3179
2.2489 | -82.319
2.2493 | -82.3215
2.2496 | -82.3245
2.2499 | -82.3258
2.2503 | -82.3273
2.2506 | -82.3306
2.251 | -82.3336
2.2513 | -82.3374
2.2516 | -82.3404
2.252 | -82.3417
2.2523 | -82.3439
2.2527 | -82.3454
2.253 | -82.3462
2.2534 | -82.3473
2.2537 | -82.3482
2.254 | -82.349
2.2544 | -82.346
2.2547 | -82.3451
2.2551 | -82.3468
2.2554 | -82.3443
2.2557 | -82.3436
2.2561 | -82.3458
2.2564 | -82.3486
2.2568 | -82.3513
2.2571 | -82.3538
2.2574 | -82.355
2.2578 | -82.3566
2.2581 | -82.3593
2.2585 | -82.3621
2.2588 | -82.3649
2.2591 | -82.3684
2.2595 | -82.3746
2.2598 | -82.3799
2.2602 | -82.3843
2.2605 | -82.3876
2.2608 | -82.3914
2.2612 | -82.3934
2.2615 | -82.394
2.2619 | -82.3973
2.2622 | -82.3999
2.2625 | -82.4036
2.2629 | -82.412
2.2632 | -82.4166
2.2636 | -82.4199
2.2639 | -82.4233
2.2642 | -82.428
2.2646 | -82.4308
2.2649 | -82.4337
2.2653 | -82.4381
2.2656 | -82.4423
2.2659 | -82.446
2.2663 | -82.4495
2.2666 | -82.4547
2.267 | -82.4608
2.2673 | -82.4663
2.2677 | -82.4702
2.268 | -82.4736
2.2683 | -82.4754
2.2687 | -82.479
2.269 | -82.4843
2.2694 | -82.4889
2.2697 | -82.4929
2.27 | -82.4966
2.2704 | -82.4988
2.2707 | -82.5014
2.2711 | -82.5057
2.2714 | -82.5091
2.2717 | -82.5128
2.2721 | -82.5144
2.2724 | -82.5188
2.2728 | -82.5237
2.2731 | -82.5288
2.2734 | -82.5337
2.2738 | -82.5376
2.2741 | -82.5404
2.2745 | -82.5441
2.2748 | -82.5478
2.2751 | -82.551
2.2755 | -82.5564
2.2758 | -82.5626
2.2762 | -82.567
2.2765 | -82.571
2.2768 | -82.5746
2.2772 | -82.5803
2.2775 | -82.5857
2.2779 | -82.5886
2.2782 | -82.5918
2.2785 | -82.596
2.2789 | -82.6015
2.2792 | -82.6042
2.2796 | -82.6075
2.2799 | -82.6125
2.2802 | -82.6164
2.2806 | -82.6203
2.2809 | -82.6246
2.2813 | -82.6302
2.2816 | -82.6356
2.282 | -82.6416
2.2823 | -82.6482
2.2826 | -82.6526
2.283 | -82.6559
2.2833 | -82.661
2.2837 | -82.665
2.284 | -82.6687
2.2843 | -82.6728
2.2847 | -82.6787
2.285 | -82.6858
2.2854 | -82.6909
2.2857 | -82.6951
2.286 | -82.6995
2.2864 | -82.7034
2.2867 | -82.708
2.2871 | -82.7145
2.2874 | -82.7201
2.2877 | -82.724
2.2881 | -82.7289
2.2884 | -82.7355
2.2888 | -82.7403
2.2891 | -82.744
2.2894 | -82.7483
2.2898 | -82.7524
2.2901 | -82.7554
2.2905 | -82.7596
2.2908 | -82.7657
2.2911 | -82.7702
2.2915 | -82.7747
2.2918 | -82.7795
2.2922 | -82.7827
2.2925 | -82.7866
2.2928 | -82.7912
2.2932 | -82.7959
2.2935 | -82.8017
2.2939 | -82.808
2.2942 | -82.811
2.2945 | -82.8161
2.2949 | -82.8199
2.2952 | -82.8253
2.2956 | -82.8307
2.2959 | -82.835
2.2963 | -82.8391
2.2966 | -82.8434
2.2969 | -82.846
2.2973 | -82.8489
2.2976 | -82.8552
2.298 | -82.8598
2.2983 | -82.8626
2.2986 | -82.8659
2.299 | -82.8699
2.2993 | -82.8707
2.2997 | -82.8741
2.3 | -82.8798
2.3003 | -82.885
2.3007 | -82.8872
2.301 | -82.8895
2.3014 | -82.8926
2.3017 | -82.8978
2.302 | -82.9035
2.3024 | -82.9074
2.3027 | -82.9113
2.3031 | -82.9165
2.3034 | -82.9222
2.3037 | -82.9273
2.3041 | -82.932
2.3044 | -82.9355
2.3048 | -82.9391
2.3051 | -82.9427
2.3054 | -82.9452
2.3058 | -82.9505
2.3061 | -82.9539
2.3065 | -82.9575
2.3068 | -82.9627
2.3071 | -82.9686
2.3075 | -82.9719
2.3078 | -82.9759
2.3082 | -82.9804
2.3085 | -82.98
2.3089 | -82.9827
2.3092 | -82.9859
2.3095 | -82.9882
2.3099 | -82.9919
2.3102 | -82.998
2.3106 | -83.0082
2.3109 | -83.0149
2.3112 | -83.0202
2.3116 | -83.0245
2.3119 | -83.0277
2.3123 | -83.0313
2.3126 | -83.035
2.3129 | -83.0378
2.3133 | -83.0402
2.3136 | -83.0426
2.314 | -83.0449
2.3143 | -83.0446
2.3146 | -83.0451
2.315 | -83.0524
2.3153 | -83.059
2.3157 | -83.0621
2.316 | -83.0651
2.3163 | -83.0661
2.3167 | -83.0672
2.317 | -83.0691
2.3174 | -83.0714
2.3177 | -83.0738
2.318 | -83.0791
2.3184 | -83.0834
2.3187 | -83.0867
2.3191 | -83.0869
2.3194 | -83.089
2.3197 | -83.093
2.3201 | -83.096
2.3204 | -83.0991
2.3208 | -83.1038
2.3211 | -83.1075
2.3214 | -83.1109
2.3218 | -83.1141
2.3221 | -83.1169
2.3225 | -83.1206
2.3228 | -83.1235
2.3232 | -83.127
2.3235 | -83.1286
2.3238 | -83.1296
2.3242 | -83.1349
2.3245 | -83.1381
2.3249 | -83.1414
2.3252 | -83.1442
2.3255 | -83.1441
2.3259 | -83.1433
2.3262 | -83.1456
2.3266 | -83.148
2.3269 | -83.15
2.3272 | -83.1509
2.3276 | -83.1522
2.3279 | -83.1511
2.3283 | -83.1536
2.3286 | -83.1564
2.3289 | -83.161
2.3293 | -83.167
2.3296 | -83.1701
2.33 | -83.1749
2.3303 | -83.1796
2.3306 | -83.1821
2.331 | -83.1847
2.3313 | -83.1884
2.3317 | -83.1931
2.332 | -83.1973
2.3323 | -83.1976
2.3327 | -83.1986
2.333 | -83.2032
2.3334 | -83.209
2.3337 | -83.212
2.334 | -83.2128
2.3344 | -83.2152
2.3347 | -83.221
2.3351 | -83.2275
2.3354 | -83.2273
2.3357 | -83.2295
2.3361 | -83.233
2.3364 | -83.2358
2.3368 | -83.2378
2.3371 | -83.239
2.3375 | -83.2417
2.3378 | -83.2408
2.3381 | -83.2415
2.3385 | -83.2418
2.3388 | -83.2455
2.3392 | -83.2495
2.3395 | -83.2526
2.3398 | -83.2545
2.3402 | -83.2576
2.3405 | -83.2594
2.3409 | -83.262
2.3412 | -83.2661
2.3415 | -83.2688
2.3419 | -83.2719
2.3422 | -83.2746
2.3426 | -83.2765
2.3429 | -83.2788
2.3432 | -83.2801
2.3436 | -83.282
2.3439 | -83.2827
2.3443 | -83.2824
2.3446 | -83.2849
2.3449 | -83.2879
2.3453 | -83.2904
2.3456 | -83.2955
2.346 | -83.2987
2.3463 | -83.3002
2.3466 | -83.3023
2.347 | -83.3043
2.3473 | -83.3067
2.3477 | -83.305
2.348 | -83.3032
2.3483 | -83.3001
2.3487 | -83.2983
2.349 | -83.3009
2.3494 | -83.3035
2.3497 | -83.3025
2.35 | -83.3032
2.3504 | -83.3056
2.3507 | -83.3077
2.3511 | -83.3079
2.3514 | -83.3115
2.3518 | -83.3136
2.3521 | -83.3135
2.3524 | -83.3175
2.3528 | -83.3186
2.3531 | -83.3205
2.3535 | -83.3216
2.3538 | -83.3232
2.3541 | -83.3254
2.3545 | -83.3266
2.3548 | -83.3274
2.3552 | -83.3296
2.3555 | -83.3291
2.3558 | -83.3277
2.3562 | -83.332
2.3565 | -83.3361
2.3569 | -83.3369
2.3572 | -83.3373
2.3575 | -83.3367
2.3579 | -83.3363
2.3582 | -83.336
2.3586 | -83.3352
2.3589 | -83.3369
2.3592 | -83.3393
2.3596 | -83.3407
2.3599 | -83.3431
2.3603 | -83.3456
2.3606 | -83.3451
2.3609 | -83.3452
2.3613 | -83.3445
2.3616 | -83.3434
2.362 | -83.3438
2.3623 | -83.3437
2.3626 | -83.3444
2.363 | -83.345
2.3633 | -83.3444
2.3637 | -83.3453
2.364 | -83.3455
2.3643 | -83.3455
2.3647 | -83.3462
2.365 | -83.3452
2.3654 | -83.3401
2.3657 | -83.3386
2.3661 | -83.3385
2.3664 | -83.3387
2.3667 | -83.339
2.3671 | -83.34
2.3674 | -83.3402
2.3678 | -83.3408
2.3681 | -83.3408
2.3684 | -83.3423
2.3688 | -83.3428
2.3691 | -83.3415
2.3695 | -83.3411
2.3698 | -83.3406
2.3701 | -83.3413
2.3705 | -83.3406
2.3708 | -83.3381
2.3712 | -83.3372
2.3715 | -83.3376
2.3718 | -83.3378
2.3722 | -83.3387
2.3725 | -83.3382
2.3729 | -83.3366
2.3732 | -83.3376
2.3735 | -83.3373
2.3739 | -83.3362
2.3742 | -83.3339
2.3746 | -83.3329
2.3749 | -83.3323
2.3752 | -83.3335
2.3756 | -83.3338
2.3759 | -83.3326
2.3763 | -83.3353
2.3766 | -83.3371
2.3769 | -83.3383
2.3773 | -83.3362
2.3776 | -83.3358
2.378 | -83.3347
2.3783 | -83.3334
2.3786 | -83.3311
2.379 | -83.3302
2.3793 | -83.329
2.3797 | -83.3273
2.38 | -83.3266
2.3804 | -83.3248
2.3807 | -83.3231
2.381 | -83.3211
2.3814 | -83.3196
2.3817 | -83.3198
2.3821 | -83.3191
2.3824 | -83.317
2.3827 | -83.316
2.3831 | -83.3161
2.3834 | -83.3159
2.3838 | -83.3142
2.3841 | -83.3135
2.3844 | -83.3107
2.3848 | -83.3075
2.3851 | -83.3053
2.3855 | -83.304
2.3858 | -83.3026
2.3861 | -83.301
2.3865 | -83.3
2.3868 | -83.3002
2.3872 | -83.2973
2.3875 | -83.2954
2.3878 | -83.2958
2.3882 | -83.2949
2.3885 | -83.2942
2.3889 | -83.2907
2.3892 | -83.2901
2.3895 | -83.288
2.3899 | -83.2861
2.3902 | -83.2827
2.3906 | -83.281
2.3909 | -83.2798
2.3912 | -83.2792
2.3916 | -83.2766
2.3919 | -83.2749
2.3923 | -83.2735
2.3926 | -83.2709
2.393 | -83.269
2.3933 | -83.2668
2.3936 | -83.2666
2.394 | -83.2663
2.3943 | -83.2615
2.3947 | -83.2593
2.395 | -83.2597
2.3953 | -83.2586
2.3957 | -83.257
2.396 | -83.2551
2.3964 | -83.253
2.3967 | -83.2527
2.397 | -83.2514
2.3974 | -83.2498
2.3977 | -83.249
2.3981 | -83.2476
2.3984 | -83.2458
2.3987 | -83.2447
2.3991 | -83.243
2.3994 | -83.2415
2.3998 | -83.2421
2.4001 | -83.2425
2.4004 | -83.2418
2.4008 | -83.2415
2.4011 | -83.2396
2.4015 | -83.2364
2.4018 | -83.2351
2.4021 | -83.2346
2.4025 | -83.2327
2.4028 | -83.2307
2.4032 | -83.2292
2.4035 | -83.2274
2.4038 | -83.2255
2.4042 | -83.2242
2.4045 | -83.2229
2.4049 | -83.2205
2.4052 | -83.2154
2.4055 | -83.2101
2.4059 | -83.2069
2.4062 | -83.2048
2.4066 | -83.2043
2.4069 | -83.2029
2.4073 | -83.1996
2.4076 | -83.1966
2.4079 | -83.1949
2.4083 | -83.1929
2.4086 | -83.1911
2.409 | -83.1888
2.4093 | -83.1864
2.4096 | -83.1863
2.41 | -83.187
2.4103 | -83.1861
2.4107 | -83.1843
2.411 | -83.1843
2.4113 | -83.1832
2.4117 | -83.1832
2.412 | -83.1823
2.4124 | -83.1816
2.4127 | -83.1803
2.413 | -83.1782
2.4134 | -83.1775
2.4137 | -83.1753
2.4141 | -83.174
2.4144 | -83.1735
2.4147 | -83.1722
2.4151 | -83.172
2.4154 | -83.1727
2.4158 | -83.1726
2.4161 | -83.1734
2.4164 | -83.1737
2.4168 | -83.1754
2.4171 | -83.1806
2.4175 | -83.1826
2.4178 | -83.1827
2.4181 | -83.1867
2.4185 | -83.1878
2.4188 | -83.1864
2.4192 | -83.1846
2.4195 | -83.1831
2.4198 | -83.1823
2.4202 | -83.1825
2.4205 | -83.1821
2.4209 | -83.1804
2.4212 | -83.1789
2.4216 | -83.1774
2.4219 | -83.1765
2.4222 | -83.1743
2.4226 | -83.1713
2.4229 | -83.1707
2.4233 | -83.1702
2.4236 | -83.1706
2.4239 | -83.1709
2.4243 | -83.1701
2.4246 | -83.1694
2.425 | -83.1666
2.4253 | -83.1651
2.4256 | -83.1644
2.426 | -83.1634
2.4263 | -83.165
2.4267 | -83.1642
2.427 | -83.1639
2.4273 | -83.1638
2.4277 | -83.1632
2.428 | -83.1637
2.4284 | -83.1639
2.4287 | -83.1628
2.429 | -83.1624
2.4294 | -83.1622
2.4297 | -83.1614
2.4301 | -83.1623
2.4304 | -83.1639
2.4307 | -83.1636
2.4311 | -83.1625
2.4314 | -83.1626
2.4318 | -83.1643
2.4321 | -83.1645
2.4324 | -83.1649
2.4328 | -83.1655
2.4331 | -83.1666
2.4335 | -83.1683
2.4338 | -83.1693
2.4341 | -83.1697
2.4345 | -83.1702
2.4348 | -83.168
2.4352 | -83.1658
2.4355 | -83.1655
2.4359 | -83.1676
2.4362 | -83.1693
2.4365 | -83.17
2.4369 | -83.1694
2.4372 | -83.1699
2.4376 | -83.1697
2.4379 | -83.1687
2.4382 | -83.169
2.4386 | -83.1699
2.4389 | -83.1714
2.4393 | -83.172
2.4396 | -83.1731
2.4399 | -83.1741
2.4403 | -83.1749
2.4406 | -83.1756
2.441 | -83.177
2.4413 | -83.1776
2.4416 | -83.1796
2.442 | -83.1813
2.4423 | -83.1804
2.4427 | -83.1801
2.443 | -83.181
2.4433 | -83.1806
2.4437 | -83.1817
2.444 | -83.1837
2.4444 | -83.1846
2.4447 | -83.1867
2.445 | -83.1876
2.4454 | -83.1878
2.4457 | -83.1878
2.4461 | -83.1879
2.4464 | -83.1879
2.4467 | -83.1877
2.4471 | -83.1881
2.4474 | -83.1879
2.4478 | -83.1879
2.4481 | -83.1894
2.4484 | -83.1907
2.4488 | -83.1914
2.4491 | -83.1922
2.4495 | -83.1934
2.4498 | -83.1959
2.4502 | -83.1973
2.4505 | -83.1975
2.4508 | -83.1978
2.4512 | -83.1998
2.4515 | -83.2003
2.4519 | -83.2015
2.4522 | -83.2026
2.4525 | -83.204
2.4529 | -83.2062
2.4532 | -83.2083
2.4536 | -83.2096
2.4539 | -83.2093
2.4542 | -83.2109
2.4546 | -83.2113
2.4549 | -83.2132
2.4553 | -83.2149
2.4556 | -83.2151
2.4559 | -83.2165
2.4563 | -83.2179
2.4566 | -83.2171
2.457 | -83.2128
2.4573 | -83.2124
2.4576 | -83.2163
2.458 | -83.2192
2.4583 | -83.218
2.4587 | -83.2151
2.459 | -83.2141
2.4593 | -83.2127
2.4597 | -83.2098
2.46 | -83.2053
2.4604 | -83.2042
2.4607 | -83.2051
2.461 | -83.2028
2.4614 | -83.2036
2.4617 | -83.2031
2.4621 | -83.1967
2.4624 | -83.1924
2.4628 | -83.1932
2.4631 | -83.1928
2.4634 | -83.1921
2.4638 | -83.1931
2.4641 | -83.1953
2.4645 | -83.1934
2.4648 | -83.1898
2.4651 | -83.1925
2.4655 | -83.1928
2.4658 | -83.1941
2.4662 | -83.1944
2.4665 | -83.1944
2.4668 | -83.1965
2.4672 | -83.1976
2.4675 | -83.1972
2.4679 | -83.1974
2.4682 | -83.1981
2.4685 | -83.1977
2.4689 | -83.197
2.4692 | -83.1974
2.4696 | -83.1967
2.4699 | -83.1966
2.4702 | -83.1973
2.4706 | -83.1953
2.4709 | -83.1924
2.4713 | -83.1923
2.4716 | -83.1948
2.4719 | -83.1946
2.4723 | -83.1959
2.4726 | -83.1962
2.473 | -83.1957
2.4733 | -83.1956
2.4736 | -83.1976
2.474 | -83.1981
2.4743 | -83.1982
2.4747 | -83.1995
2.475 | -83.2019
2.4753 | -83.1995
2.4757 | -83.2
2.476 | -83.1994
2.4764 | -83.1986
2.4767 | -83.1982
2.4771 | -83.1977
2.4774 | -83.1981
2.4777 | -83.1996
2.4781 | -83.1988
2.4784 | -83.1965
2.4788 | -83.1946
2.4791 | -83.1929
2.4794 | -83.1915
2.4798 | -83.1904
2.4801 | -83.1891
2.4805 | -83.188
2.4808 | -83.1879
2.4811 | -83.1901
2.4815 | -83.1898
2.4818 | -83.1883
2.4822 | -83.1869
2.4825 | -83.1859
2.4828 | -83.1849
2.4832 | -83.1855
2.4835 | -83.1851
2.4839 | -83.1835
2.4842 | -83.1823
2.4845 | -83.1815
2.4849 | -83.1793
2.4852 | -83.1762
2.4856 | -83.1749
2.4859 | -83.1741
2.4862 | -83.1758
2.4866 | -83.1785
2.4869 | -83.1776
2.4873 | -83.1763
2.4876 | -83.1716
2.4879 | -83.1681
2.4883 | -83.1667
2.4886 | -83.1652
2.489 | -83.1634
2.4893 | -83.1589
2.4896 | -83.1548
2.49 | -83.1528
2.4903 | -83.1498
2.4907 | -83.1485
2.491 | -83.1475
2.4914 | -83.147
2.4917 | -83.1454
2.492 | -83.1434
2.4924 | -83.1387
2.4927 | -83.1346
2.4931 | -83.1329
2.4934 | -83.1307
2.4937 | -83.1283
2.4941 | -83.1258
2.4944 | -83.1237
2.4948 | -83.1201
2.4951 | -83.1164
2.4954 | -83.1145
2.4958 | -83.114
2.4961 | -83.1113
2.4965 | -83.1097
2.4968 | -83.1068
2.4971 | -83.1052
2.4975 | -83.1018
2.4978 | -83.098
2.4982 | -83.0938
2.4985 | -83.09
2.4988 | -83.0832
2.4992 | -83.0741
2.4995 | -83.0701
2.4999 | -83.0673
2.5002 | -83.0644
2.5005 | -83.0613
2.5009 | -83.059
2.5012 | -83.0559
2.5016 | -83.0531
2.5019 | -83.0508
2.5022 | -83.0461
2.5026 | -83.0375
2.5029 | -83.0322
2.5033 | -83.0273
2.5036 | -83.0243
2.504 | -83.0209
2.5043 | -83.0168
2.5046 | -83.0141
2.505 | -83.0091
2.5053 | -83.0062
2.5057 | -83.0036
2.506 | -82.9978
2.5063 | -82.994
2.5067 | -82.9928
2.507 | -82.9895
2.5074 | -82.9847
2.5077 | -82.9794
2.508 | -82.975
2.5084 | -82.9709
2.5087 | -82.9663
2.5091 | -82.9623
2.5094 | -82.9595
2.5097 | -82.9557
2.5101 | -82.9514
2.5104 | -82.948
2.5108 | -82.944
2.5111 | -82.94
2.5114 | -82.9365
2.5118 | -82.9327
2.5121 | -82.9292
2.5125 | -82.9209
2.5128 | -82.9137
2.5131 | -82.9098
2.5135 | -82.9031
2.5138 | -82.8975
2.5142 | -82.8927
2.5145 | -82.8875
2.5148 | -82.8821
2.5152 | -82.8782
2.5155 | -82.8732
2.5159 | -82.8689
2.5162 | -82.8642
2.5165 | -82.8602
2.5169 | -82.8555
2.5172 | -82.8512
2.5176 | -82.8497
2.5179 | -82.8489
2.5183 | -82.8472
2.5186 | -82.8432
2.5189 | -82.8393
2.5193 | -82.8355
2.5196 | -82.8294
2.52 | -82.8233
2.5203 | -82.8164
2.5206 | -82.8113
2.521 | -82.8063
2.5213 | -82.802
2.5217 | -82.8003
2.522 | -82.7959
2.5223 | -82.7887
2.5227 | -82.7835
2.523 | -82.7787
2.5234 | -82.7726
2.5237 | -82.7656
2.524 | -82.7607
2.5244 | -82.7564
2.5247 | -82.7525
2.5251 | -82.7474
2.5254 | -82.744
2.5257 | -82.7418
2.5261 | -82.7376
2.5264 | -82.7314
2.5268 | -82.7233
2.5271 | -82.7176
2.5274 | -82.7128
2.5278 | -82.7095
2.5281 | -82.7054
2.5285 | -82.7017
2.5288 | -82.6983
2.5291 | -82.694
2.5295 | -82.6894
2.5298 | -82.6848
2.5302 | -82.6807
2.5305 | -82.6759
2.5308 | -82.6716
2.5312 | -82.6677
2.5315 | -82.6646
2.5319 | -82.6612
2.5322 | -82.6578
2.5326 | -82.6532
2.5329 | -82.6485
2.5332 | -82.6418
2.5336 | -82.6348
2.5339 | -82.6296
2.5343 | -82.6265
2.5346 | -82.6221
2.5349 | -82.6174
2.5353 | -82.6109
2.5356 | -82.6045
2.536 | -82.6013
2.5363 | -82.5977
2.5366 | -82.5934
2.537 | -82.5885
2.5373 | -82.5836
2.5377 | -82.5789
2.538 | -82.5729
2.5383 | -82.5661
2.5387 | -82.5615
2.539 | -82.556
2.5394 | -82.5515
2.5397 | -82.5472
2.54 | -82.5432
2.5404 | -82.5385
2.5407 | -82.535
2.5411 | -82.531
2.5414 | -82.5273
2.5417 | -82.5237
2.5421 | -82.5167
2.5424 | -82.5086
2.5428 | -82.5035
2.5431 | -82.4995
2.5434 | -82.4962
2.5438 | -82.4928
2.5441 | -82.4904
2.5445 | -82.4858
2.5448 | -82.4828
2.5451 | -82.4797
2.5455 | -82.477
2.5458 | -82.4737
2.5462 | -82.4669
2.5465 | -82.4583
2.5469 | -82.4529
2.5472 | -82.4492
2.5475 | -82.4462
2.5479 | -82.4412
2.5482 | -82.4375
2.5486 | -82.4341
2.5489 | -82.4294
2.5492 | -82.4244
2.5496 | -82.4185
2.5499 | -82.4117
2.5503 | -82.4084
2.5506 | -82.4057
2.5509 | -82.4017
2.5513 | -82.3984
2.5516 | -82.3943
2.552 | -82.3906
2.5523 | -82.3861
2.5526 | -82.3812
2.553 | -82.3776
2.5533 | -82.3753
2.5537 | -82.3727
2.554 | -82.3704
2.5543 | -82.3685
2.5547 | -82.3652
2.555 | -82.3599
2.5554 | -82.3548
2.5557 | -82.3517
2.556 | -82.3481
2.5564 | -82.3457
2.5567 | -82.3432
2.5571 | -82.3387
2.5574 | -82.3344
2.5577 | -82.33
2.5581 | -82.3265
2.5584 | -82.3234
2.5588 | -82.3185
2.5591 | -82.3151
2.5594 | -82.3121
2.5598 | -82.3072
2.5601 | -82.302
2.5605 | -82.2983
2.5608 | -82.2944
2.5612 | -82.2925
2.5615 | -82.2894
2.5618 | -82.2869
2.5622 | -82.2849
2.5625 | -82.2826
2.5629 | -82.279
2.5632 | -82.2757
2.5635 | -82.271
2.5639 | -82.2694
2.5642 | -82.2674
2.5646 | -82.2637
2.5649 | -82.2587
2.5652 | -82.254
2.5656 | -82.2517
2.5659 | -82.2491
2.5663 | -82.2449
2.5666 | -82.2386
2.5669 | -82.2362
2.5673 | -82.2336
2.5676 | -82.231
2.568 | -82.2268
2.5683 | -82.2241
2.5686 | -82.2211
2.569 | -82.2159
2.5693 | -82.2107
2.5697 | -82.2071
2.57 | -82.2043
2.5703 | -82.2005
2.5707 | -82.1968
2.571 | -82.194
2.5714 | -82.1924
2.5717 | -82.1897
2.572 | -82.1871
2.5724 | -82.1847
2.5727 | -82.1814
2.5731 | -82.1728
2.5734 | -82.1681
2.5737 | -82.1679
2.5741 | -82.1645
2.5744 | -82.1604
2.5748 | -82.1574
2.5751 | -82.1544
2.5755 | -82.1522
2.5758 | -82.1493
2.5761 | -82.1473
2.5765 | -82.1423
2.5768 | -82.1392
2.5772 | -82.1368
2.5775 | -82.1347
2.5778 | -82.1318
2.5782 | -82.129
2.5785 | -82.1288
2.5789 | -82.1293
2.5792 | -82.1286
2.5795 | -82.1267
2.5799 | -82.1236
2.5802 | -82.1171
2.5806 | -82.1125
2.5809 | -82.1116
2.5812 | -82.1108
2.5816 | -82.1076
2.5819 | -82.1038
2.5823 | -82.1021
2.5826 | -82.0999
2.5829 | -82.0978
2.5833 | -82.0946
2.5836 | -82.0927
2.584 | -82.0887
2.5843 | -82.0872
2.5846 | -82.0853
2.585 | -82.0835
2.5853 | -82.081
2.5857 | -82.079
2.586 | -82.0776
2.5863 | -82.0751
2.5867 | -82.0713
2.587 | -82.0661
2.5874 | -82.0593
2.5877 | -82.0556
2.588 | -82.0539
2.5884 | -82.0511
2.5887 | -82.0483
2.5891 | -82.046
2.5894 | -82.0426
2.5898 | -82.0408
2.5901 | -82.039
2.5904 | -82.0367
2.5908 | -82.0326
2.5911 | -82.0286
2.5915 | -82.025
2.5918 | -82.022
2.5921 | -82.0201
2.5925 | -82.0171
2.5928 | -82.016
2.5932 | -82.0112
2.5935 | -82.0062
2.5938 | -82.0037
2.5942 | -82.0001
2.5945 | -81.998
2.5949 | -81.9963
2.5952 | -81.9946
2.5955 | -81.9933
2.5959 | -81.9919
2.5962 | -81.9903
2.5966 | -81.9881
2.5969 | -81.9863
2.5972 | -81.9853
2.5976 | -81.9828
2.5979 | -81.981
2.5983 | -81.9779
2.5986 | -81.9761
2.5989 | -81.9748
2.5993 | -81.9746
2.5996 | -81.9741
2.6 | -81.9733
2.6003 | -81.9714
2.6006 | -81.9692
2.601 | -81.9642
2.6013 | -81.961
2.6017 | -81.9589
2.602 | -81.9571
2.6024 | -81.955
2.6027 | -81.9536
2.603 | -81.953
2.6034 | -81.951
2.6037 | -81.9496
2.6041 | -81.9484
2.6044 | -81.9467
2.6047 | -81.9451
2.6051 | -81.9442
2.6054 | -81.9426
2.6058 | -81.9411
2.6061 | -81.9393
2.6064 | -81.9383
2.6068 | -81.9369
2.6071 | -81.9328
2.6075 | -81.9281
2.6078 | -81.9258
2.6081 | -81.9227
2.6085 | -81.919
2.6088 | -81.9167
2.6092 | -81.9144
2.6095 | -81.913
2.6098 | -81.912
2.6102 | -81.9116
2.6105 | -81.9092
2.6109 | -81.9044
2.6112 | -81.9026
2.6115 | -81.9012
2.6119 | -81.8979
2.6122 | -81.8956
2.6126 | -81.8945
2.6129 | -81.8909
2.6132 | -81.8852
2.6136 | -81.8818
2.6139 | -81.8811
2.6143 | -81.8804
2.6146 | -81.8781
2.6149 | -81.875
2.6153 | -81.8735
2.6156 | -81.872
2.616 | -81.8699
2.6163 | -81.8678
2.6167 | -81.8652
2.617 | -81.8629
2.6173 | -81.8611
2.6177 | -81.8596
2.618 | -81.8578
2.6184 | -81.8575
2.6187 | -81.856
2.619 | -81.8542
2.6194 | -81.8532
2.6197 | -81.8523
2.6201 | -81.8509
2.6204 | -81.8495
2.6207 | -81.8461
2.6211 | -81.8415
2.6214 | -81.8374
2.6218 | -81.8349
2.6221 | -81.8336
2.6224 | -81.8327
2.6228 | -81.8301
2.6231 | -81.8257
2.6235 | -81.8223
2.6238 | -81.8213
2.6241 | -81.8207
2.6245 | -81.8206
2.6248 | -81.8187
2.6252 | -81.8167
2.6255 | -81.8146
2.6258 | -81.8129
2.6262 | -81.8111
2.6265 | -81.8092
2.6269 | -81.8065
2.6272 | -81.8033
2.6275 | -81.8013
2.6279 | -81.7995
2.6282 | -81.798
2.6286 | -81.7946
2.6289 | -81.7922
2.6292 | -81.7905
2.6296 | -81.7887
2.6299 | -81.7863
2.6303 | -81.785
2.6306 | -81.7846
2.631 | -81.7827
2.6313 | -81.7807
2.6316 | -81.7793
2.632 | -81.776
2.6323 | -81.7709
2.6327 | -81.7692
2.633 | -81.7679
2.6333 | -81.7665
2.6337 | -81.7648
2.634 | -81.7646
2.6344 | -81.7651
2.6347 | -81.7621
2.635 | -81.7566
2.6354 | -81.7518
2.6357 | -81.7518
2.6361 | -81.7527
2.6364 | -81.752
2.6367 | -81.7513
2.6371 | -81.7518
2.6374 | -81.751
2.6378 | -81.7482
2.6381 | -81.7456
2.6384 | -81.7454
2.6388 | -81.7427
2.6391 | -81.7408
2.6395 | -81.7434
2.6398 | -81.7452
2.6401 | -81.7435
2.6405 | -81.7418
2.6408 | -81.7397
2.6412 | -81.7366
2.6415 | -81.733
2.6418 | -81.7308
2.6422 | -81.7298
2.6425 | -81.7253
2.6429 | -81.717
2.6432 | -81.7127
2.6435 | -81.71
2.6439 | -81.7076
2.6442 | -81.705
2.6446 | -81.7022
2.6449 | -81.6994
2.6453 | -81.6974
2.6456 | -81.6941
2.6459 | -81.6895
2.6463 | -81.6871
2.6466 | -81.6852
2.647 | -81.6833
2.6473 | -81.6789
2.6476 | -81.6752
2.648 | -81.6721
2.6483 | -81.668
2.6487 | -81.6649
2.649 | -81.6641
2.6493 | -81.6618
2.6497 | -81.6593
2.65 | -81.6573
2.6504 | -81.6547
2.6507 | -81.6527
2.651 | -81.6502
2.6514 | -81.6477
2.6517 | -81.6461
2.6521 | -81.6438
2.6524 | -81.6429
2.6527 | -81.6404
2.6531 | -81.6369
2.6534 | -81.6353
2.6538 | -81.6338
2.6541 | -81.6301
2.6544 | -81.6263
2.6548 | -81.624
2.6551 | -81.6232
2.6555 | -81.621
2.6558 | -81.6188
2.6561 | -81.6164
2.6565 | -81.615
2.6568 | -81.6132
2.6572 | -81.6125
2.6575 | -81.611
2.6578 | -81.6092
2.6582 | -81.6083
2.6585 | -81.6051
2.6589 | -81.6016
2.6592 | -81.5999
2.6596 | -81.5978
2.6599 | -81.5962
2.6602 | -81.5959
2.6606 | -81.5937
2.6609 | -81.5916
2.6613 | -81.5894
2.6616 | -81.5874
2.6619 | -81.5858
2.6623 | -81.5856
2.6626 | -81.5844
2.663 | -81.5839
2.6633 | -81.5827
2.6636 | -81.5811
2.664 | -81.5802
2.6643 | -81.5787
2.6647 | -81.5781
2.665 | -81.5766
2.6653 | -81.5753
2.6657 | -81.5741
2.666 | -81.5721
2.6664 | -81.5701
2.6667 | -81.568
2.667 | -81.5683
2.6674 | -81.5682
2.6677 | -81.5664
2.6681 | -81.5647
2.6684 | -81.5641
2.6687 | -81.565
2.6691 | -81.5648
2.6694 | -81.5632
2.6698 | -81.5626
2.6701 | -81.5613
2.6704 | -81.5613
2.6708 | -81.5598
2.6711 | -81.5568
2.6715 | -81.5552
2.6718 | -81.5554
2.6722 | -81.5558
2.6725 | -81.5533
2.6728 | -81.5538
2.6732 | -81.555
2.6735 | -81.555
2.6739 | -81.5535
2.6742 | -81.5514
2.6745 | -81.5505
2.6749 | -81.55
2.6752 | -81.5488
2.6756 | -81.5469
2.6759 | -81.5466
2.6762 | -81.5452
2.6766 | -81.5441
2.6769 | -81.5433
2.6773 | -81.5425
2.6776 | -81.5405
2.6779 | -81.5384
2.6783 | -81.5387
2.6786 | -81.538
2.679 | -81.5395
2.6793 | -81.5426
2.6796 | -81.5379
2.68 | -81.5363
2.6803 | -81.5353
2.6807 | -81.5365
2.681 | -81.5373
2.6813 | -81.5366
2.6817 | -81.5367
2.682 | -81.5355
2.6824 | -81.5334
2.6827 | -81.5328
2.683 | -81.5346
2.6834 | -81.5337
2.6837 | -81.5342
2.6841 | -81.5338
2.6844 | -81.5317
2.6847 | -81.5314
2.6851 | -81.5283
2.6854 | -81.5278
2.6858 | -81.527
2.6861 | -81.5267
2.6865 | -81.524
2.6868 | -81.5223
2.6871 | -81.5224
2.6875 | -81.5221
2.6878 | -81.5229
2.6882 | -81.5218
2.6885 | -81.5209
2.6888 | -81.5203
2.6892 | -81.519
2.6895 | -81.5164
2.6899 | -81.5149
2.6902 | -81.5127
2.6905 | -81.5122
2.6909 | -81.5119
2.6912 | -81.5109
2.6916 | -81.512
2.6919 | -81.5116
2.6922 | -81.5112
2.6926 | -81.5103
2.6929 | -81.5068
2.6933 | -81.5064
2.6936 | -81.5089
2.6939 | -81.5105
2.6943 | -81.5083
2.6946 | -81.5087
2.695 | -81.5085
2.6953 | -81.503
2.6956 | -81.4987
2.696 | -81.4981
2.6963 | -81.4977
2.6967 | -81.4976
2.697 | -81.4976
2.6973 | -81.4956
2.6977 | -81.4938
2.698 | -81.4934
2.6984 | -81.4926
2.6987 | -81.4937
2.699 | -81.4935
2.6994 | -81.4933
2.6997 | -81.4926
2.7001 | -81.4929
2.7004 | -81.4914
2.7008 | -81.4912
2.7011 | -81.4912
2.7014 | -81.4875
2.7018 | -81.4873
2.7021 | -81.488
2.7025 | -81.4877
2.7028 | -81.4877
2.7031 | -81.4876
2.7035 | -81.4826
2.7038 | -81.4776
2.7042 | -81.476
2.7045 | -81.4737
2.7048 | -81.473
2.7052 | -81.472
2.7055 | -81.4728
2.7059 | -81.4717
2.7062 | -81.4706
2.7065 | -81.4697
2.7069 | -81.4692
2.7072 | -81.4683
2.7076 | -81.4661
2.7079 | -81.4644
2.7082 | -81.4627
2.7086 | -81.462
2.7089 | -81.4636
2.7093 | -81.4601
2.7096 | -81.4559
2.7099 | -81.4544
2.7103 | -81.4525
2.7106 | -81.4504
2.711 | -81.4482
2.7113 | -81.4482
2.7116 | -81.4476
2.712 | -81.4466
2.7123 | -81.4457
2.7127 | -81.4447
2.713 | -81.4431
2.7134 | -81.4417
2.7137 | -81.4404
2.714 | -81.4385
2.7144 | -81.4362
2.7147 | -81.4332
2.7151 | -81.4318
2.7154 | -81.4306
2.7157 | -81.4307
2.7161 | -81.4322
2.7164 | -81.4317
2.7168 | -81.4299
2.7171 | -81.4284
2.7174 | -81.4279
2.7178 | -81.4277
2.7181 | -81.4285
2.7185 | -81.4297
2.7188 | -81.4303
2.7191 | -81.4291
2.7195 | -81.4282
2.7198 | -81.4275
2.7202 | -81.4267
2.7205 | -81.4259
2.7208 | -81.4261
2.7212 | -81.4255
2.7215 | -81.4241
2.7219 | -81.4228
2.7222 | -81.4217
2.7225 | -81.4198
2.7229 | -81.419
2.7232 | -81.4184
2.7236 | -81.4173
2.7239 | -81.4153
2.7242 | -81.4146
2.7246 | -81.4136
2.7249 | -81.4132
2.7253 | -81.4142
2.7256 | -81.4148
2.7259 | -81.4142
2.7263 | -81.4121
2.7266 | -81.4101
2.727 | -81.4092
2.7273 | -81.4082
2.7277 | -81.406
2.728 | -81.4052
2.7283 | -81.4045
2.7287 | -81.403
2.729 | -81.4023
2.7294 | -81.4005
2.7297 | -81.4009
2.73 | -81.4031
2.7304 | -81.4012
2.7307 | -81.3987
2.7311 | -81.3978
2.7314 | -81.3965
2.7317 | -81.3949
2.7321 | -81.3923
2.7324 | -81.3919
2.7328 | -81.39
2.7331 | -81.3891
2.7334 | -81.3895
2.7338 | -81.3906
2.7341 | -81.3886
2.7345 | -81.3862
2.7348 | -81.3841
2.7351 | -81.3803
2.7355 | -81.3792
2.7358 | -81.3772
2.7362 | -81.3747
2.7365 | -81.3747
2.7368 | -81.3739
2.7372 | -81.3726
2.7375 | -81.3718
2.7379 | -81.3708
2.7382 | -81.3691
2.7385 | -81.3679
2.7389 | -81.3667
2.7392 | -81.3656
2.7396 | -81.365
2.7399 | -81.3639
2.7402 | -81.3628
2.7406 | -81.3631
2.7409 | -81.3622
2.7413 | -81.3615
2.7416 | -81.3615
2.742 | -81.3603
2.7423 | -81.3595
2.7426 | -81.358
2.743 | -81.3555
2.7433 | -81.3552
2.7437 | -81.3551
2.744 | -81.3547
2.7443 | -81.3545
2.7447 | -81.3548
2.745 | -81.3535
2.7454 | -81.3515
2.7457 | -81.3495
2.746 | -81.3464
2.7464 | -81.3446
2.7467 | -81.342
2.7471 | -81.3406
2.7474 | -81.3395
2.7477 | -81.3383
2.7481 | -81.338
2.7484 | -81.3378
2.7488 | -81.3344
2.7491 | -81.3329
2.7494 | -81.332
2.7498 | -81.3317
2.7501 | -81.331
2.7505 | -81.3312
2.7508 | -81.3297
2.7511 | -81.325
2.7515 | -81.3229
2.7518 | -81.3246
2.7522 | -81.3248
2.7525 | -81.3246
2.7528 | -81.3222
2.7532 | -81.3208
2.7535 | -81.3191
2.7539 | -81.3173
2.7542 | -81.3182
2.7545 | -81.318
2.7549 | -81.3164
2.7552 | -81.3142
2.7556 | -81.3125
2.7559 | -81.3115
2.7563 | -81.3102
2.7566 | -81.3082
2.7569 | -81.3065
2.7573 | -81.3055
2.7576 | -81.3046
2.758 | -81.3038
2.7583 | -81.3063
2.7586 | -81.307
2.759 | -81.3039
2.7593 | -81.3016
2.7597 | -81.3013
2.76 | -81.3025
2.7603 | -81.3029
2.7607 | -81.3017
2.761 | -81.2995
2.7614 | -81.2982
2.7617 | -81.2973
2.762 | -81.2963
2.7624 | -81.2969
2.7627 | -81.2949
2.7631 | -81.2933
2.7634 | -81.292
2.7637 | -81.2909
2.7641 | -81.29
2.7644 | -81.2874
2.7648 | -81.2847
2.7651 | -81.2835
2.7654 | -81.2825
2.7658 | -81.2817
2.7661 | -81.2828
2.7665 | -81.2806
2.7668 | -81.278
2.7671 | -81.2772
2.7675 | -81.2764
2.7678 | -81.2739
2.7682 | -81.2719
2.7685 | -81.2709
2.7688 | -81.2693
2.7692 | -81.2686
2.7695 | -81.2683
2.7699 | -81.2673
2.7702 | -81.2669
2.7706 | -81.2672
2.7709 | -81.2665
2.7712 | -81.2654
2.7716 | -81.2632
2.7719 | -81.2622
2.7723 | -81.2612
2.7726 | -81.2601
2.7729 | -81.2597
2.7733 | -81.2582
2.7736 | -81.2592
2.774 | -81.262
2.7743 | -81.2625
2.7746 | -81.2598
2.775 | -81.257
2.7753 | -81.2537
2.7757 | -81.2528
2.776 | -81.2519
2.7763 | -81.2506
2.7767 | -81.2492
2.777 | -81.2471
2.7774 | -81.248
2.7777 | -81.2477
2.778 | -81.2451
2.7784 | -81.2431
2.7787 | -81.2393
2.7791 | -81.2378
2.7794 | -81.2367
2.7797 | -81.2385
2.7801 | -81.2392
2.7804 | -81.2405
2.7808 | -81.2388
2.7811 | -81.2384
2.7814 | -81.2386
2.7818 | -81.2369
2.7821 | -81.2359
2.7825 | -81.2348
2.7828 | -81.2339
2.7832 | -81.2332
2.7835 | -81.2325
2.7838 | -81.2319
2.7842 | -81.232
2.7845 | -81.2319
2.7849 | -81.2316
2.7852 | -81.2302
2.7855 | -81.2303
2.7859 | -81.23
2.7862 | -81.2298
2.7866 | -81.2285
2.7869 | -81.2267
2.7872 | -81.2249
2.7876 | -81.2241
2.7879 | -81.2243
2.7883 | -81.2257
2.7886 | -81.2262
2.7889 | -81.224
2.7893 | -81.223
2.7896 | -81.2229
2.79 | -81.2234
2.7903 | -81.2196
2.7906 | -81.2176
2.791 | -81.2177
2.7913 | -81.2159
2.7917 | -81.2128
2.792 | -81.2079
2.7923 | -81.2027
2.7927 | -81.2002
2.793 | -81.1986
2.7934 | -81.1982
2.7937 | -81.198
2.794 | -81.1979
2.7944 | -81.1961
2.7947 | -81.1936
2.7951 | -81.1917
2.7954 | -81.188
2.7957 | -81.1855
2.7961 | -81.1845
2.7964 | -81.1841
2.7968 | -81.1839
2.7971 | -81.1817
2.7974 | -81.1811
2.7978 | -81.1776
2.7981 | -81.1765
2.7985 | -81.1776
2.7988 | -81.1778
2.7992 | -81.1758
2.7995 | -81.1762
2.7998 | -81.1744
2.8002 | -81.1734
2.8005 | -81.172
2.8009 | -81.17
2.8012 | -81.1685
2.8015 | -81.1676
2.8019 | -81.1692
2.8022 | -81.169
2.8026 | -81.1669
2.8029 | -81.1652
2.8032 | -81.1624
2.8036 | -81.1608
2.8039 | -81.1601
2.8043 | -81.1587
2.8046 | -81.1579
2.8049 | -81.1567
2.8053 | -81.1554
2.8056 | -81.1534
2.806 | -81.1526
2.8063 | -81.1508
2.8066 | -81.1499
2.807 | -81.1492
2.8073 | -81.1473
2.8077 | -81.1452
2.808 | -81.1444
2.8083 | -81.142
2.8087 | -81.1394
2.809 | -81.1377
2.8094 | -81.1365
2.8097 | -81.1365
2.81 | -81.1369
2.8104 | -81.1372
2.8107 | -81.1371
2.8111 | -81.136
2.8114 | -81.1356
2.8118 | -81.1349
2.8121 | -81.1331
2.8124 | -81.135
2.8128 | -81.1354
2.8131 | -81.1333
2.8135 | -81.1309
2.8138 | -81.1274
2.8141 | -81.1245
2.8145 | -81.1211
2.8148 | -81.1199
2.8152 | -81.1173
2.8155 | -81.1123
2.8158 | -81.1078
2.8162 | -81.105
2.8165 | -81.1028
2.8169 | -81.1014
2.8172 | -81.0997
2.8175 | -81.0973
2.8179 | -81.0962
2.8182 | -81.0934
2.8186 | -81.0914
2.8189 | -81.0888
2.8192 | -81.0866
2.8196 | -81.0848
2.8199 | -81.0823
2.8203 | -81.0799
2.8206 | -81.079
2.8209 | -81.0775
2.8213 | -81.0752
2.8216 | -81.0751
2.822 | -81.0764
2.8223 | -81.0759
2.8226 | -81.0737
2.823 | -81.0713
2.8233 | -81.069
2.8237 | -81.066
2.824 | -81.0642
2.8243 | -81.0636
2.8247 | -81.0637
2.825 | -81.0634
2.8254 | -81.0614
2.8257 | -81.0557
2.8261 | -81.0513
2.8264 | -81.0488
2.8267 | -81.047
2.8271 | -81.0453
2.8274 | -81.0442
2.8278 | -81.044
2.8281 | -81.0421
2.8284 | -81.0398
2.8288 | -81.038
2.8291 | -81.0347
2.8295 | -81.0321
2.8298 | -81.0273
2.8301 | -81.0247
2.8305 | -81.0221
2.8308 | -81.0193
2.8312 | -81.0157
2.8315 | -81.0116
2.8318 | -81.0093
2.8322 | -81.0056
2.8325 | -81.0034
2.8329 | -81.0013
2.8332 | -81.0002
2.8335 | -81.0019
2.8339 | -81.0006
2.8342 | -80.9982
2.8346 | -80.9944
2.8349 | -80.9878
2.8352 | -80.9843
2.8356 | -80.9815
2.8359 | -80.9781
2.8363 | -80.9758
2.8366 | -80.9738
2.8369 | -80.9703
2.8373 | -80.967
2.8376 | -80.9633
2.838 | -80.9601
2.8383 | -80.9584
2.8386 | -80.9555
2.839 | -80.9513
2.8393 | -80.9486
2.8397 | -80.9463
2.84 | -80.9411
2.8404 | -80.9363
2.8407 | -80.9324
2.841 | -80.9292
2.8414 | -80.9251
2.8417 | -80.9194
2.8421 | -80.9148
2.8424 | -80.9113
2.8427 | -80.908
2.8431 | -80.9044
2.8434 | -80.9005
2.8438 | -80.8961
2.8441 | -80.8921
2.8444 | -80.8889
2.8448 | -80.8838
2.8451 | -80.8803
2.8455 | -80.8785
2.8458 | -80.8765
2.8461 | -80.8736
2.8465 | -80.8701
2.8468 | -80.8666
2.8472 | -80.8623
2.8475 | -80.8567
2.8478 | -80.8501
2.8482 | -80.844
2.8485 | -80.8395
2.8489 | -80.8343
2.8492 | -80.8307
2.8495 | -80.8278
2.8499 | -80.8217
2.8502 | -80.8173
2.8506 | -80.8156
2.8509 | -80.8109
2.8512 | -80.8069
2.8516 | -80.8033
2.8519 | -80.7996
2.8523 | -80.7961
2.8526 | -80.7929
2.8529 | -80.7895
2.8533 | -80.7864
2.8536 | -80.783
2.854 | -80.7789
2.8543 | -80.7747
2.8547 | -80.7713
2.855 | -80.7663
2.8553 | -80.7619
2.8557 | -80.7579
2.856 | -80.7539
2.8564 | -80.7496
2.8567 | -80.7443
2.857 | -80.7368
2.8574 | -80.7323
2.8577 | -80.7282
2.8581 | -80.7246
2.8584 | -80.722
2.8587 | -80.7187
2.8591 | -80.7142
2.8594 | -80.7095
2.8598 | -80.7056
2.8601 | -80.7022
2.8604 | -80.6984
2.8608 | -80.6935
2.8611 | -80.6896
2.8615 | -80.6841
2.8618 | -80.6784
2.8621 | -80.6748
2.8625 | -80.6708
2.8628 | -80.666
2.8632 | -80.6621
2.8635 | -80.6575
2.8638 | -80.6531
2.8642 | -80.649
2.8645 | -80.6453
2.8649 | -80.6418
2.8652 | -80.6403
2.8655 | -80.6348
2.8659 | -80.6292
2.8662 | -80.6255
2.8666 | -80.622
2.8669 | -80.6174
2.8672 | -80.6139
2.8676 | -80.6107
2.8679 | -80.6069
2.8683 | -80.6048
2.8686 | -80.6019
2.869 | -80.5978
2.8693 | -80.5945
2.8696 | -80.5912
2.87 | -80.588
2.8703 | -80.5834
2.8707 | -80.5799
2.871 | -80.5753
2.8713 | -80.5697
2.8717 | -80.5638
2.872 | -80.5597
2.8724 | -80.5565
2.8727 | -80.5554
2.873 | -80.5495
2.8734 | -80.5458
2.8737 | -80.5413
2.8741 | -80.536
2.8744 | -80.5309
2.8747 | -80.5272
2.8751 | -80.5238
2.8754 | -80.5196
2.8758 | -80.5174
2.8761 | -80.515
2.8764 | -80.5122
2.8768 | -80.5087
2.8771 | -80.5052
2.8775 | -80.501
2.8778 | -80.4976
2.8781 | -80.4935
2.8785 | -80.4898
2.8788 | -80.4868
2.8792 | -80.4837
2.8795 | -80.4791
2.8798 | -80.4756
2.8802 | -80.4734
2.8805 | -80.4698
2.8809 | -80.466
2.8812 | -80.4629
2.8816 | -80.4601
2.8819 | -80.4578
2.8822 | -80.4549
2.8826 | -80.4514
2.8829 | -80.4475
2.8833 | -80.4458
2.8836 | -80.4423
2.8839 | -80.4388
2.8843 | -80.4341
2.8846 | -80.4291
2.885 | -80.426
2.8853 | -80.4227
2.8856 | -80.42
2.886 | -80.4166
2.8863 | -80.4126
2.8867 | -80.4087
2.887 | -80.4054
2.8873 | -80.4022
2.8877 | -80.3966
2.888 | -80.3948
2.8884 | -80.3918
2.8887 | -80.3901
2.889 | -80.3865
2.8894 | -80.38
2.8897 | -80.3782
2.8901 | -80.3784
2.8904 | -80.3743
2.8907 | -80.3699
2.8911 | -80.3659
2.8914 | -80.3613
2.8918 | -80.359
2.8921 | -80.3572
2.8924 | -80.3544
2.8928 | -80.3515
2.8931 | -80.3486
2.8935 | -80.3457
2.8938 | -80.343
2.8941 | -80.3407
2.8945 | -80.3372
2.8948 | -80.334
2.8952 | -80.3315
2.8955 | -80.3293
2.8959 | -80.3273
2.8962 | -80.3271
2.8965 | -80.325
2.8969 | -80.3228
2.8972 | -80.3208
2.8976 | -80.3185
2.8979 | -80.3153
2.8982 | -80.3128
2.8986 | -80.3138
2.8989 | -80.3144
2.8993 | -80.3131
2.8996 | -80.3111
2.8999 | -80.3062
2.9003 | -80.3025
2.9006 | -80.2987
2.901 | -80.2962
2.9013 | -80.2932
2.9016 | -80.2896
2.902 | -80.2867
2.9023 | -80.2851
2.9027 | -80.2836
2.903 | -80.282
2.9033 | -80.2808
2.9037 | -80.2793
2.904 | -80.277
2.9044 | -80.2777
2.9047 | -80.278
2.905 | -80.2755
2.9054 | -80.2729
2.9057 | -80.2712
2.9061 | -80.2702
2.9064 | -80.2686
2.9067 | -80.2658
2.9071 | -80.2625
2.9074 | -80.2608
2.9078 | -80.261
2.9081 | -80.2589
2.9085 | -80.2568
2.9088 | -80.2557
2.9091 | -80.2549
2.9095 | -80.2535
2.9098 | -80.2522
2.9102 | -80.2502
2.9105 | -80.2492
2.9108 | -80.2481
2.9112 | -80.2439
2.9115 | -80.2382
2.9119 | -80.2355
2.9122 | -80.2341
2.9125 | -80.2321
2.9129 | -80.2303
2.9132 | -80.2287
2.9136 | -80.2267
2.9139 | -80.2242
2.9142 | -80.2223
2.9146 | -80.2214
2.9149 | -80.2201
2.9153 | -80.217
2.9156 | -80.2148
2.9159 | -80.2134
2.9163 | -80.2113
2.9166 | -80.2098
2.917 | -80.2085
2.9173 | -80.2074
2.9176 | -80.1993
2.918 | -80.198
2.9183 | -80.1964
2.9187 | -80.1948
2.919 | -80.1942
2.9193 | -80.1932
2.9197 | -80.1921
2.92 | -80.1913
2.9204 | -80.19
2.9207 | -80.1873
2.921 | -80.1857
2.9214 | -80.1843
2.9217 | -80.1832
2.9221 | -80.1829
2.9224 | -80.1813
2.9228 | -80.181
2.9231 | -80.1808
2.9234 | -80.1796
2.9238 | -80.1772
2.9241 | -80.1767
2.9245 | -80.1757
2.9248 | -80.1748
2.9251 | -80.1722
2.9255 | -80.1692
2.9258 | -80.1684
2.9262 | -80.1679
2.9265 | -80.1672
2.9268 | -80.1677
2.9272 | -80.1676
2.9275 | -80.1673
2.9279 | -80.1669
2.9282 | -80.1662
2.9285 | -80.1658
2.9289 | -80.1657
2.9292 | -80.1654
2.9296 | -80.1638
2.9299 | -80.1627
2.9302 | -80.1625
2.9306 | -80.1619
2.9309 | -80.1613
2.9313 | -80.1618
2.9316 | -80.1621
2.9319 | -80.1616
2.9323 | -80.1621
2.9326 | -80.161
2.933 | -80.1605
2.9333 | -80.1595
2.9336 | -80.1587
2.934 | -80.158
2.9343 | -80.157
2.9347 | -80.1567
2.935 | -80.1564
2.9353 | -80.1558
2.9357 | -80.1545
2.936 | -80.1534
2.9364 | -80.1537
2.9367 | -80.1532
2.9371 | -80.1544
2.9374 | -80.154
2.9377 | -80.1537
2.9381 | -80.1534
2.9384 | -80.1531
2.9388 | -80.1527
2.9391 | -80.1519
2.9394 | -80.1506
2.9398 | -80.1507
2.9401 | -80.1498
2.9405 | -80.149
2.9408 | -80.149
2.9411 | -80.1479
2.9415 | -80.1472
2.9418 | -80.1469
2.9422 | -80.1475
2.9425 | -80.1475
2.9428 | -80.1477
2.9432 | -80.1488
2.9435 | -80.1496
2.9439 | -80.1512
2.9442 | -80.152
2.9445 | -80.153
2.9449 | -80.1533
2.9452 | -80.1532
2.9456 | -80.1528
2.9459 | -80.1534
2.9462 | -80.1534
2.9466 | -80.1541
2.9469 | -80.1533
2.9473 | -80.1518
2.9476 | -80.152
2.9479 | -80.1522
2.9483 | -80.1527
2.9486 | -80.1526
2.949 | -80.1527
2.9493 | -80.1535
2.9496 | -80.1558
2.95 | -80.1572
2.9503 | -80.1574
2.9507 | -80.1577
2.951 | -80.1582
2.9514 | -80.1583
2.9517 | -80.1582
2.952 | -80.1582
2.9524 | -80.1593
2.9527 | -80.1596
2.9531 | -80.1607
2.9534 | -80.1614
2.9537 | -80.1613
2.9541 | -80.1624
2.9544 | -80.1631
2.9548 | -80.1632
2.9551 | -80.1622
2.9554 | -80.1621
2.9558 | -80.1626
2.9561 | -80.1628
2.9565 | -80.1629
2.9568 | -80.1652
2.9571 | -80.1664
2.9575 | -80.1671
2.9578 | -80.1691
2.9582 | -80.1698
2.9585 | -80.1686
2.9588 | -80.167
2.9592 | -80.1675
2.9595 | -80.166
2.9599 | -80.1654
2.9602 | -80.1665
2.9605 | -80.1661
2.9609 | -80.1639
2.9612 | -80.163
2.9616 | -80.1639
2.9619 | -80.1653
2.9622 | -80.1663
2.9626 | -80.1672
2.9629 | -80.1682
2.9633 | -80.1688
2.9636 | -80.1709
2.9639 | -80.1722
2.9643 | -80.1729
2.9646 | -80.173
2.965 | -80.1738
2.9653 | -80.1754
2.9657 | -80.1775
2.966 | -80.1788
2.9663 | -80.1796
2.9667 | -80.1814
2.967 | -80.183
2.9674 | -80.1857
2.9677 | -80.183
2.968 | -80.1804
2.9684 | -80.1808
2.9687 | -80.1804
2.9691 | -80.1804
2.9694 | -80.1817
2.9697 | -80.1835
2.9701 | -80.1833
2.9704 | -80.1825
2.9708 | -80.1829
2.9711 | -80.1836
2.9714 | -80.1851
2.9718 | -80.1856
2.9721 | -80.1863
2.9725 | -80.1877
2.9728 | -80.1884
2.9731 | -80.1902
2.9735 | -80.1907
2.9738 | -80.19
2.9742 | -80.1908
2.9745 | -80.1926
2.9748 | -80.1933
2.9752 | -80.1923
2.9755 | -80.1918
2.9759 | -80.1932
2.9762 | -80.1942
2.9765 | -80.1964
2.9769 | -80.1975
2.9772 | -80.1996
2.9776 | -80.2
2.9779 | -80.2009
2.9783 | -80.2035
2.9786 | -80.2058
2.9789 | -80.2066
2.9793 | -80.2045
2.9796 | -80.2048
2.98 | -80.2056
2.9803 | -80.2062
2.9806 | -80.2058
2.981 | -80.2063
2.9813 | -80.2061
2.9817 | -80.2066
2.982 | -80.2068
2.9823 | -80.2081
2.9827 | -80.2091
2.983 | -80.2091
2.9834 | -80.2096
2.9837 | -80.2101
2.984 | -80.2085
2.9844 | -80.2083
2.9847 | -80.208
2.9851 | -80.209
2.9854 | -80.2096
2.9857 | -80.2115
2.9861 | -80.2123
2.9864 | -80.214
2.9868 | -80.2151
2.9871 | -80.2164
2.9874 | -80.2179
2.9878 | -80.2196
2.9881 | -80.2212
2.9885 | -80.2229
2.9888 | -80.2234
2.9891 | -80.2244
2.9895 | -80.2257
2.9898 | -80.2276
2.9902 | -80.2284
2.9905 | -80.2287
2.9908 | -80.2299
2.9912 | -80.2314
2.9915 | -80.2352
2.9919 | -80.2376
2.9922 | -80.2387
2.9926 | -80.2406
2.9929 | -80.2431
2.9932 | -80.2469
2.9936 | -80.2514
2.9939 | -80.2512
2.9943 | -80.2511
2.9946 | -80.252
2.9949 | -80.253
2.9953 | -80.2545
2.9956 | -80.2564
2.996 | -80.2578
2.9963 | -80.2585
2.9966 | -80.2619
2.997 | -80.267
2.9973 | -80.2689
2.9977 | -80.2701
2.998 | -80.2718
2.9983 | -80.2729
2.9987 | -80.2755
2.999 | -80.2764
2.9994 | -80.2797
2.9997 | -80.281
3 | -80.2827
3.0004 | -80.285
3.0007 | -80.2866
3.0011 | -80.2857
3.0014 | -80.2864
3.0017 | -80.2883
3.0021 | -80.2904
3.0024 | -80.291
3.0028 | -80.2919
3.0031 | -80.2937
3.0034 | -80.296
3.0038 | -80.2962
3.0041 | -80.2975
3.0045 | -80.2995
3.0048 | -80.3002
3.0051 | -80.3012
3.0055 | -80.3019
3.0058 | -80.3027
3.0062 | -80.3028
3.0065 | -80.3028
3.0069 | -80.3026
3.0072 | -80.3036
3.0075 | -80.3054
3.0079 | -80.3073
3.0082 | -80.3081
3.0086 | -80.3089
3.0089 | -80.3099
3.0092 | -80.311
3.0096 | -80.3129
3.0099 | -80.3151
3.0103 | -80.3163
3.0106 | -80.3185
3.0109 | -80.321
3.0113 | -80.3225
3.0116 | -80.3235
3.012 | -80.325
3.0123 | -80.3258
3.0126 | -80.3276
3.013 | -80.3296
3.0133 | -80.3308
3.0137 | -80.3316
3.014 | -80.3326
3.0143 | -80.3348
3.0147 | -80.3368
3.015 | -80.338
3.0154 | -80.3395
3.0157 | -80.3406
3.016 | -80.3427
3.0164 | -80.3449
3.0167 | -80.3477
3.0171 | -80.35
3.0174 | -80.3525
3.0177 | -80.3541
3.0181 | -80.3557
3.0184 | -80.3597
3.0188 | -80.3636
3.0191 | -80.3665
3.0194 | -80.3692
3.0198 | -80.3696
3.0201 | -80.3703
3.0205 | -80.3724
3.0208 | -80.3741
3.0212 | -80.3765
3.0215 | -80.3778
3.0218 | -80.3788
3.0222 | -80.3795
3.0225 | -80.3814
3.0229 | -80.3832
3.0232 | -80.3853
3.0235 | -80.3874
3.0239 | -80.3872
3.0242 | -80.3861
3.0246 | -80.386
3.0249 | -80.3867
3.0252 | -80.3883
3.0256 | -80.3905
3.0259 | -80.3912
3.0263 | -80.3928
3.0266 | -80.3947
3.0269 | -80.3968
3.0273 | -80.3992
3.0276 | -80.4004
3.028 | -80.4018
3.0283 | -80.4026
3.0286 | -80.4052
3.029 | -80.4076
3.0293 | -80.4089
3.0297 | -80.4092
3.03 | -80.4096
3.0303 | -80.4104
3.0307 | -80.4117
3.031 | -80.413
3.0314 | -80.4133
3.0317 | -80.4137
3.032 | -80.415
3.0324 | -80.4163
3.0327 | -80.4182
3.0331 | -80.4203
3.0334 | -80.4222
3.0337 | -80.4256
3.0341 | -80.4265
3.0344 | -80.4284
3.0348 | -80.4295
3.0351 | -80.4313
3.0355 | -80.4336
3.0358 | -80.4368
3.0361 | -80.4389
3.0365 | -80.4421
3.0368 | -80.4429
3.0372 | -80.4435
3.0375 | -80.4458
3.0378 | -80.4477
3.0382 | -80.4487
3.0385 | -80.4497
3.0389 | -80.4513
3.0392 | -80.4551
3.0395 | -80.4568
3.0399 | -80.4575
3.0402 | -80.4584
3.0406 | -80.4589
3.0409 | -80.4602
3.0412 | -80.4612
3.0416 | -80.4614
3.0419 | -80.4635
3.0423 | -80.4655
3.0426 | -80.466
3.0429 | -80.4674
3.0433 | -80.4693
3.0436 | -80.4713
3.044 | -80.4735
3.0443 | -80.475
3.0446 | -80.4778
3.045 | -80.4805
3.0453 | -80.4811
3.0457 | -80.4832
3.046 | -80.4837
3.0463 | -80.4841
3.0467 | -80.4833
3.047 | -80.4829
3.0474 | -80.4839
3.0477 | -80.484
3.048 | -80.4852
3.0484 | -80.4874
3.0487 | -80.4886
3.0491 | -80.4907
3.0494 | -80.4925
3.0498 | -80.4948
3.0501 | -80.4965
3.0504 | -80.4978
3.0508 | -80.4982
3.0511 | -80.4974
3.0515 | -80.4966
3.0518 | -80.4965
3.0521 | -80.4987
3.0525 | -80.5013
3.0528 | -80.5022
3.0532 | -80.5039
3.0535 | -80.5053
3.0538 | -80.5074
3.0542 | -80.5093
3.0545 | -80.5106
3.0549 | -80.5118
3.0552 | -80.514
3.0555 | -80.517
3.0559 | -80.5182
3.0562 | -80.5193
3.0566 | -80.5217
3.0569 | -80.5232
3.0572 | -80.5245
3.0576 | -80.5255
3.0579 | -80.5268
3.0583 | -80.5276
3.0586 | -80.5285
3.0589 | -80.5298
3.0593 | -80.5309
3.0596 | -80.5318
3.06 | -80.5331
3.0603 | -80.5344
3.0606 | -80.5365
3.061 | -80.5378
3.0613 | -80.5393
3.0617 | -80.541
3.062 | -80.5433
3.0623 | -80.5459
3.0627 | -80.5504
3.063 | -80.5509
3.0634 | -80.5522
3.0637 | -80.5543
3.0641 | -80.5558
3.0644 | -80.5575
3.0647 | -80.5607
3.0651 | -80.562
3.0654 | -80.564
3.0658 | -80.5656
3.0661 | -80.5663
3.0664 | -80.5677
3.0668 | -80.5681
3.0671 | -80.5706
3.0675 | -80.5718
3.0678 | -80.5751
3.0681 | -80.5786
3.0685 | -80.579
3.0688 | -80.5804
3.0692 | -80.5817
3.0695 | -80.5812
3.0698 | -80.5836
3.0702 | -80.5843
3.0705 | -80.5851
3.0709 | -80.5862
3.0712 | -80.5873
3.0715 | -80.5887
3.0719 | -80.5914
3.0722 | -80.5937
3.0726 | -80.5933
3.0729 | -80.5941
3.0732 | -80.5976
3.0736 | -80.5996
3.0739 | -80.6008
3.0743 | -80.6017
3.0746 | -80.6016
3.0749 | -80.6031
3.0753 | -80.6044
3.0756 | -80.6042
3.076 | -80.6048
3.0763 | -80.606
3.0767 | -80.6064
3.077 | -80.6063
3.0773 | -80.6085
3.0777 | -80.611
3.078 | -80.615
3.0784 | -80.6181
3.0787 | -80.6221
3.079 | -80.6244
3.0794 | -80.626
3.0797 | -80.6283
3.0801 | -80.6324
3.0804 | -80.6339
3.0807 | -80.6359
3.0811 | -80.6387
3.0814 | -80.6413
3.0818 | -80.6429
3.0821 | -80.6451
3.0824 | -80.6472
3.0828 | -80.6489
3.0831 | -80.6505
3.0835 | -80.6514
3.0838 | -80.6532
3.0841 | -80.6546
3.0845 | -80.6556
3.0848 | -80.6581
3.0852 | -80.6599
3.0855 | -80.6622
3.0858 | -80.6654
3.0862 | -80.6675
3.0865 | -80.6705
3.0869 | -80.6715
3.0872 | -80.6734
3.0875 | -80.6752
3.0879 | -80.6775
3.0882 | -80.68
3.0886 | -80.6805
3.0889 | -80.6821
3.0892 | -80.6838
3.0896 | -80.6862
3.0899 | -80.6886
3.0903 | -80.691
3.0906 | -80.6934
3.091 | -80.6957
3.0913 | -80.6985
3.0916 | -80.701
3.092 | -80.7037
3.0923 | -80.7068
3.0927 | -80.708
3.093 | -80.7092
3.0933 | -80.7109
3.0937 | -80.7131
3.094 | -80.716
3.0944 | -80.7183
3.0947 | -80.7211
3.095 | -80.7247
3.0954 | -80.7262
3.0957 | -80.73
3.0961 | -80.7326
3.0964 | -80.7345
3.0967 | -80.7367
3.0971 | -80.7393
3.0974 | -80.742
3.0978 | -80.7472
3.0981 | -80.7508
3.0984 | -80.7517
3.0988 | -80.7548
3.0991 | -80.7578
3.0995 | -80.7603
3.0998 | -80.7638
3.1001 | -80.7655
3.1005 | -80.7654
3.1008 | -80.7666
3.1012 | -80.7694
3.1015 | -80.7707
3.1018 | -80.7719
3.1022 | -80.7739
3.1025 | -80.7757
3.1029 | -80.7779
3.1032 | -80.7798
3.1035 | -80.7817
3.1039 | -80.7835
3.1042 | -80.7868
3.1046 | -80.788
3.1049 | -80.7893
3.1053 | -80.7912
3.1056 | -80.7932
3.1059 | -80.7953
3.1063 | -80.797
3.1066 | -80.7984
3.107 | -80.8025
3.1073 | -80.8043
3.1076 | -80.8059
3.108 | -80.8081
3.1083 | -80.8098
3.1087 | -80.8106
3.109 | -80.8141
3.1093 | -80.8181
3.1097 | -80.8217
3.11 | -80.8244
3.1104 | -80.8276
3.1107 | -80.8308
3.111 | -80.8338
3.1114 | -80.8353
3.1117 | -80.837
3.1121 | -80.8384
3.1124 | -80.8396
3.1127 | -80.8408
3.1131 | -80.8435
3.1134 | -80.844
3.1138 | -80.8443
3.1141 | -80.8474
3.1144 | -80.8491
3.1148 | -80.8514
3.1151 | -80.8537
3.1155 | -80.8561
3.1158 | -80.857
3.1161 | -80.8585
3.1165 | -80.8608
3.1168 | -80.8627
3.1172 | -80.8649
3.1175 | -80.8666
3.1178 | -80.8688
3.1182 | -80.8714
3.1185 | -80.872
3.1189 | -80.8734
3.1192 | -80.8753
3.1196 | -80.8776
3.1199 | -80.8804
3.1202 | -80.8815
3.1206 | -80.8839
3.1209 | -80.8889
3.1213 | -80.8959
3.1216 | -80.8992
3.1219 | -80.9007
3.1223 | -80.9036
3.1226 | -80.9066
3.123 | -80.9087
3.1233 | -80.9111
3.1236 | -80.9142
3.124 | -80.9169
3.1243 | -80.9194
3.1247 | -80.9204
3.125 | -80.9226
3.1253 | -80.9261
3.1257 | -80.9291
3.126 | -80.9315
3.1264 | -80.9346
3.1267 | -80.9365
3.127 | -80.9384
3.1274 | -80.9405
3.1277 | -80.9458
3.1281 | -80.9537
3.1284 | -80.9574
3.1287 | -80.9598
3.1291 | -80.9628
3.1294 | -80.9644
3.1298 | -80.9655
3.1301 | -80.9657
3.1304 | -80.9674
3.1308 | -80.9708
3.1311 | -80.9743
3.1315 | -80.9753
3.1318 | -80.9775
3.1322 | -80.98
3.1325 | -80.9819
3.1328 | -80.9842
3.1332 | -80.9857
3.1335 | -80.9882
3.1339 | -80.9902
3.1342 | -80.9926
3.1345 | -80.9947
3.1349 | -80.9956
3.1352 | -80.9966
3.1356 | -80.9974
3.1359 | -81.0013
3.1362 | -81.0052
3.1366 | -81.0071
3.1369 | -81.009
3.1373 | -81.0103
3.1376 | -81.0123
3.1379 | -81.0137
3.1383 | -81.0153
3.1386 | -81.0183
3.139 | -81.0202
3.1393 | -81.0224
3.1396 | -81.0235
3.14 | -81.0257
3.1403 | -81.0276
3.1407 | -81.0306
3.141 | -81.0338
3.1413 | -81.0357
3.1417 | -81.0371
3.142 | -81.0376
3.1424 | -81.0391
3.1427 | -81.04
3.143 | -81.0417
3.1434 | -81.0418
3.1437 | -81.0438
3.1441 | -81.0462
3.1444 | -81.0491
3.1447 | -81.0533
3.1451 | -81.0556
3.1454 | -81.0576
3.1458 | -81.0561
3.1461 | -81.056
3.1465 | -81.0567
3.1468 | -81.0593
3.1471 | -81.0621
3.1475 | -81.0624
3.1478 | -81.0643
3.1482 | -81.0655
3.1485 | -81.0667
3.1488 | -81.0683
3.1492 | -81.0693
3.1495 | -81.0705
3.1499 | -81.0696
3.1502 | -81.0669
3.1505 | -81.0653
3.1509 | -81.0628
3.1512 | -81.0598
3.1516 | -81.0589
3.1519 | -81.0613
3.1522 | -81.0634
3.1526 | -81.0653
3.1529 | -81.0668
3.1533 | -81.0678
3.1536 | -81.0671
3.1539 | -81.0668
3.1543 | -81.067
3.1546 | -81.0682
3.155 | -81.0691
3.1553 | -81.0621
3.1556 | -81.0536
3.156 | -81.0512
3.1563 | -81.0462
3.1567 | -81.0436
3.157 | -81.0397
3.1573 | -81.0403
3.1577 | -81.041
3.158 | -81.0411
3.1584 | -81.0417
3.1587 | -81.0365
3.159 | -81.0332
3.1594 | -81.0353
3.1597 | -81.0352
3.1601 | -81.0225
3.1604 | -81.021
3.1608 | -81.0198
3.1611 | -81.0179
3.1614 | -81.0155
3.1618 | -81.009
3.1621 | -81.0044
3.1625 | -81.0043
3.1628 | -81.0027
3.1631 | -80.9999
3.1635 | -80.9968
3.1638 | -80.9928
3.1642 | -80.9874
3.1645 | -80.983
3.1648 | -80.9817
3.1652 | -80.9802
3.1655 | -80.9763
3.1659 | -80.9706
3.1662 | -80.9686
3.1665 | -80.9682
3.1669 | -80.9666
3.1672 | -80.9611
3.1676 | -80.9581
3.1679 | -80.9565
3.1682 | -80.9534
3.1686 | -80.9514
3.1689 | -80.9494
3.1693 | -80.9484
3.1696 | -80.9465
3.1699 | -80.9448
3.1703 | -80.9427
3.1706 | -80.939
3.171 | -80.9344
3.1713 | -80.9305
3.1716 | -80.9303
3.172 | -80.9264
3.1723 | -80.922
3.1727 | -80.9205
3.173 | -80.9141
3.1733 | -80.9139
3.1737 | -80.9111
3.174 | -80.9076
3.1744 | -80.9082
3.1747 | -80.9071
3.1751 | -80.904
3.1754 | -80.9009
3.1757 | -80.8975
3.1761 | -80.8941
3.1764 | -80.8894
3.1768 | -80.8874
3.1771 | -80.8842
3.1774 | -80.8812
3.1778 | -80.8779
3.1781 | -80.8746
3.1785 | -80.8716
3.1788 | -80.8676
3.1791 | -80.864
3.1795 | -80.8611
3.1798 | -80.8579
3.1802 | -80.8543
3.1805 | -80.849
3.1808 | -80.8444
3.1812 | -80.8405
3.1815 | -80.8402
3.1819 | -80.8391
3.1822 | -80.8361
3.1825 | -80.8333
3.1829 | -80.8273
3.1832 | -80.8237
3.1836 | -80.8196
3.1839 | -80.8151
3.1842 | -80.8121
3.1846 | -80.8096
3.1849 | -80.8058
3.1853 | -80.8017
3.1856 | -80.7977
3.1859 | -80.7929
3.1863 | -80.7886
3.1866 | -80.7842
3.187 | -80.78
3.1873 | -80.7755
3.1877 | -80.7712
3.188 | -80.7683
3.1883 | -80.7643
3.1887 | -80.7598
3.189 | -80.757
3.1894 | -80.7564
3.1897 | -80.7531
3.19 | -80.748
3.1904 | -80.7422
3.1907 | -80.7371
3.1911 | -80.7324
3.1914 | -80.7272
3.1917 | -80.7215
3.1921 | -80.717
3.1924 | -80.7122
3.1928 | -80.7027
3.1931 | -80.6974
3.1934 | -80.6923
3.1938 | -80.6891
3.1941 | -80.6876
3.1945 | -80.6826
3.1948 | -80.677
3.1951 | -80.6716
3.1955 | -80.6673
3.1958 | -80.6623
3.1962 | -80.6576
3.1965 | -80.6523
3.1968 | -80.6471
3.1972 | -80.6406
3.1975 | -80.6346
3.1979 | -80.63
3.1982 | -80.6241
3.1985 | -80.6187
3.1989 | -80.6154
3.1992 | -80.6105
3.1996 | -80.6056
3.1999 | -80.5993
3.2002 | -80.5935
3.2006 | -80.5867
3.2009 | -80.5801
3.2013 | -80.5746
3.2016 | -80.5682
3.202 | -80.5615
3.2023 | -80.5538
3.2026 | -80.5469
3.203 | -80.5408
3.2033 | -80.535
3.2037 | -80.5286
3.204 | -80.522
3.2043 | -80.5159
3.2047 | -80.5095
3.205 | -80.5032
3.2054 | -80.4977
3.2057 | -80.4919
3.206 | -80.4864
3.2064 | -80.4803
3.2067 | -80.474
3.2071 | -80.4676
3.2074 | -80.4611
3.2077 | -80.4544
3.2081 | -80.4482
3.2084 | -80.4426
3.2088 | -80.4368
3.2091 | -80.4298
3.2094 | -80.4253
3.2098 | -80.4209
3.2101 | -80.4145
3.2105 | -80.4079
3.2108 | -80.401
3.2111 | -80.3953
3.2115 | -80.3894
3.2118 | -80.3832
3.2122 | -80.3765
3.2125 | -80.3717
3.2128 | -80.3654
3.2132 | -80.3593
3.2135 | -80.3525
3.2139 | -80.3464
3.2142 | -80.3401
3.2145 | -80.3334
3.2149 | -80.3268
3.2152 | -80.3204
3.2156 | -80.3136
3.2159 | -80.3071
3.2163 | -80.3004
3.2166 | -80.2942
3.2169 | -80.2882
3.2173 | -80.2831
3.2176 | -80.2764
3.218 | -80.2697
3.2183 | -80.2631
3.2186 | -80.258
3.219 | -80.2525
3.2193 | -80.2464
3.2197 | -80.2397
3.22 | -80.2329
3.2203 | -80.2257
3.2207 | -80.2188
3.221 | -80.2132
3.2214 | -80.2085
3.2217 | -80.2037
3.222 | -80.1989
3.2224 | -80.1942
3.2227 | -80.1881
3.2231 | -80.1822
3.2234 | -80.1758
3.2237 | -80.1702
3.2241 | -80.1643
3.2244 | -80.1581
3.2248 | -80.1502
3.2251 | -80.1426
3.2254 | -80.1371
3.2258 | -80.1314
3.2261 | -80.1253
3.2265 | -80.1193
3.2268 | -80.1128
3.2271 | -80.1065
3.2275 | -80.1008
3.2278 | -80.0946
3.2282 | -80.089
3.2285 | -80.0832
3.2288 | -80.0773
3.2292 | -80.0722
3.2295 | -80.0666
3.2299 | -80.061
3.2302 | -80.055
3.2306 | -80.049
3.2309 | -80.0434
3.2312 | -80.038
3.2316 | -80.0326
3.2319 | -80.0299
3.2323 | -80.0246
3.2326 | -80.0194
3.2329 | -80.015
3.2333 | -80.0103
3.2336 | -80.0053
3.234 | -79.9999
3.2343 | -79.9944
3.2346 | -79.9901
3.235 | -79.9855
3.2353 | -79.9803
3.2357 | -79.9756
3.236 | -79.9709
3.2363 | -79.9658
3.2367 | -79.9611
3.237 | -79.9567
3.2374 | -79.9531
3.2377 | -79.9485
3.238 | -79.9438
3.2384 | -79.9397
3.2387 | -79.9356
3.2391 | -79.9315
3.2394 | -79.9281
3.2397 | -79.9246
3.2401 | -79.9207
3.2404 | -79.9167
3.2408 | -79.9127
3.2411 | -79.9088
3.2414 | -79.9059
3.2418 | -79.9048
3.2421 | -79.9025
3.2425 | -79.9001
3.2428 | -79.8972
3.2431 | -79.8936
3.2435 | -79.8901
3.2438 | -79.8869
3.2442 | -79.8838
3.2445 | -79.8803
3.2449 | -79.8771
3.2452 | -79.8737
3.2455 | -79.8705
3.2459 | -79.8677
3.2462 | -79.8648
3.2466 | -79.8623
3.2469 | -79.8584
3.2472 | -79.856
3.2476 | -79.8549
3.2479 | -79.8541
3.2483 | -79.8525
3.2486 | -79.8504
3.2489 | -79.8479
3.2493 | -79.8461
3.2496 | -79.8449
3.25 | -79.8435
3.2503 | -79.8416
3.2506 | -79.8401
3.251 | -79.8379
3.2513 | -79.8356
3.2517 | -79.8381
3.252 | -79.8353
3.2523 | -79.8335
3.2527 | -79.8327
3.253 | -79.8313
3.2534 | -79.8306
3.2537 | -79.8299
3.254 | -79.8288
3.2544 | -79.8279
3.2547 | -79.8273
3.2551 | -79.8263
3.2554 | -79.8256
3.2557 | -79.8259
3.2561 | -79.8259
3.2564 | -79.8253
3.2568 | -79.8241
3.2571 | -79.8228
3.2574 | -79.8223
3.2578 | -79.8208
3.2581 | -79.8202
3.2585 | -79.8204
3.2588 | -79.8211
3.2592 | -79.8219
3.2595 | -79.8231
3.2598 | -79.8237
3.2602 | -79.8241
3.2605 | -79.8256
3.2609 | -79.8261
3.2612 | -79.8271
3.2615 | -79.8336
3.2619 | -79.835
3.2622 | -79.8355
3.2626 | -79.8349
3.2629 | -79.8353
3.2632 | -79.8365
3.2636 | -79.8384
3.2639 | -79.8401
3.2643 | -79.8417
3.2646 | -79.8439
3.2649 | -79.8456
3.2653 | -79.8463
3.2656 | -79.8483
3.266 | -79.8512
3.2663 | -79.8561
3.2666 | -79.8606
3.267 | -79.863
3.2673 | -79.8659
3.2677 | -79.868
3.268 | -79.8701
3.2683 | -79.8732
3.2687 | -79.8755
3.269 | -79.8777
3.2694 | -79.8804
3.2697 | -79.8831
3.27 | -79.8858
3.2704 | -79.8887
3.2707 | -79.893
3.2711 | -79.9017
3.2714 | -79.9065
3.2717 | -79.9111
3.2721 | -79.9143
3.2724 | -79.9183
3.2728 | -79.9241
3.2731 | -79.9289
3.2735 | -79.9323
3.2738 | -79.9356
3.2741 | -79.9394
3.2745 | -79.9432
3.2748 | -79.947
3.2752 | -79.9513
3.2755 | -79.9555
3.2758 | -79.9597
3.2762 | -79.9652
3.2765 | -79.9696
3.2769 | -79.9742
3.2772 | -79.979
3.2775 | -79.9839
3.2779 | -79.9885
3.2782 | -79.9934
3.2786 | -79.9988
3.2789 | -80.0044
3.2792 | -80.0093
3.2796 | -80.0147
3.2799 | -80.0207
3.2803 | -80.0259
3.2806 | -80.0309
3.2809 | -80.042
3.2813 | -80.0475
3.2816 | -80.054
3.282 | -80.0602
3.2823 | -80.066
3.2826 | -80.072
3.283 | -80.0787
3.2833 | -80.0853
3.2837 | -80.093
3.284 | -80.0998
3.2843 | -80.1059
3.2847 | -80.1125
3.285 | -80.1191
3.2854 | -80.1256
3.2857 | -80.1322
3.2861 | -80.1391
3.2864 | -80.1469
3.2867 | -80.1551
3.2871 | -80.1632
3.2874 | -80.1708
3.2878 | -80.1795
3.2881 | -80.187
3.2884 | -80.1947
3.2888 | -80.2023
3.2891 | -80.2098
3.2895 | -80.2186
3.2898 | -80.2271
3.2901 | -80.2363
3.2905 | -80.2468
3.2908 | -80.2589
3.2912 | -80.267
3.2915 | -80.2767
3.2918 | -80.2861
3.2922 | -80.295
3.2925 | -80.3039
3.2929 | -80.3127
3.2932 | -80.3219
3.2935 | -80.3316
3.2939 | -80.3413
3.2942 | -80.3496
3.2946 | -80.3581
3.2949 | -80.3671
3.2952 | -80.3761
3.2956 | -80.3855
3.2959 | -80.3961
3.2963 | -80.4065
3.2966 | -80.4153
3.2969 | -80.4248
3.2973 | -80.4359
3.2976 | -80.4467
3.298 | -80.4567
3.2983 | -80.4673
3.2986 | -80.4783
3.299 | -80.4884
3.2993 | -80.4985
3.2997 | -80.5084
3.3 | -80.5186
3.3004 | -80.5285
3.3007 | -80.5421
3.301 | -80.5525
3.3014 | -80.5627
3.3017 | -80.5732
3.3021 | -80.5834
3.3024 | -80.5939
3.3027 | -80.605
3.3031 | -80.616
3.3034 | -80.627
3.3038 | -80.6379
3.3041 | -80.649
3.3044 | -80.6602
3.3048 | -80.6713
3.3051 | -80.6826
3.3055 | -80.6943
3.3058 | -80.7064
3.3061 | -80.7177
3.3065 | -80.7302
3.3068 | -80.7428
3.3072 | -80.7547
3.3075 | -80.767
3.3078 | -80.7793
3.3082 | -80.7904
3.3085 | -80.8018
3.3089 | -80.8136
3.3092 | -80.8266
3.3095 | -80.8396
3.3099 | -80.8518
3.3102 | -80.8632
3.3106 | -80.8762
3.3109 | -80.8889
3.3112 | -80.9027
3.3116 | -80.9149
3.3119 | -80.9268
3.3123 | -80.9393
3.3126 | -80.9512
3.3129 | -80.9636
3.3133 | -80.9757
3.3136 | -80.9873
3.314 | -80.9999
3.3143 | -81.0132
3.3147 | -81.0267
3.315 | -81.04
3.3153 | -81.0538
3.3157 | -81.0675
3.316 | -81.0803
3.3164 | -81.0927
3.3167 | -81.105
3.317 | -81.1185
3.3174 | -81.1323
3.3177 | -81.1459
3.3181 | -81.1597
3.3184 | -81.1736
3.3187 | -81.1873
3.3191 | -81.2004
3.3194 | -81.2141
3.3198 | -81.2279
3.3201 | -81.2417
3.3204 | -81.2552
3.3208 | -81.2691
3.3211 | -81.2833
3.3215 | -81.2979
3.3218 | -81.3119
3.3221 | -81.3253
3.3225 | -81.3391
3.3228 | -81.3521
3.3232 | -81.3648
3.3235 | -81.3777
3.3238 | -81.3924
3.3242 | -81.4061
3.3245 | -81.4194
3.3249 | -81.434
3.3252 | -81.4472
3.3255 | -81.4598
3.3259 | -81.4732
3.3262 | -81.4876
3.3266 | -81.5024
3.3269 | -81.5158
3.3272 | -81.5296
3.3276 | -81.543
3.3279 | -81.5566
3.3283 | -81.5712
3.3286 | -81.5854
3.329 | -81.5986
3.3293 | -81.6119
3.3296 | -81.6249
3.33 | -81.6347
3.3303 | -81.6486
3.3307 | -81.6635
3.331 | -81.6774
3.3313 | -81.6917
3.3317 | -81.7059
3.332 | -81.7188
3.3324 | -81.7334
3.3327 | -81.7494
3.333 | -81.7645
3.3334 | -81.7781
3.3337 | -81.7914
3.3341 | -81.8049
3.3344 | -81.8182
3.3347 | -81.8319
3.3351 | -81.8456
3.3354 | -81.8598
3.3358 | -81.8738
3.3361 | -81.8882
3.3364 | -81.9026
3.3368 | -81.9174
3.3371 | -81.9304
3.3375 | -81.9452
3.3378 | -81.9611
3.3381 | -81.9745
3.3385 | -81.9884
3.3388 | -82.0035
3.3392 | -82.0185
3.3395 | -82.0326
3.3398 | -82.0434
3.3402 | -82.0568
3.3405 | -82.071
3.3409 | -82.0859
3.3412 | -82.1013
3.3415 | -82.1147
3.3419 | -82.1284
3.3422 | -82.1422
3.3426 | -82.1542
3.3429 | -82.1669
3.3433 | -82.1822
3.3436 | -82.1961
3.3439 | -82.2099
3.3443 | -82.2234
3.3446 | -82.2379
3.345 | -82.2536
3.3453 | -82.2666
3.3456 | -82.2802
3.346 | -82.2934
3.3463 | -82.3063
3.3467 | -82.3185
3.347 | -82.3294
3.3473 | -82.341
3.3477 | -82.3531
3.348 | -82.3662
3.3484 | -82.3789
3.3487 | -82.3915
3.349 | -82.4042
3.3494 | -82.4155
3.3497 | -82.4244
3.3501 | -82.4382
3.3504 | -82.4509
3.3507 | -82.4628
3.3511 | -82.4758
3.3514 | -82.4887
3.3518 | -82.5004
3.3521 | -82.5127
3.3524 | -82.5249
3.3528 | -82.5378
3.3531 | -82.5508
3.3535 | -82.5634
3.3538 | -82.5753
3.3541 | -82.5875
3.3545 | -82.5998
3.3548 | -82.6121
3.3552 | -82.6254
3.3555 | -82.6398
3.3559 | -82.6521
3.3562 | -82.6634
3.3565 | -82.6751
3.3569 | -82.6889
3.3572 | -82.7014
3.3576 | -82.7135
3.3579 | -82.7252
3.3582 | -82.7373
3.3586 | -82.7472
3.3589 | -82.7581
3.3593 | -82.7683
3.3596 | -82.7742
3.3599 | -82.7852
3.3603 | -82.7962
3.3606 | -82.8087
3.361 | -82.8215
3.3613 | -82.8342
3.3616 | -82.8469
3.362 | -82.8586
3.3623 | -82.8698
3.3627 | -82.8811
3.363 | -82.8918
3.3633 | -82.9029
3.3637 | -82.9125
3.364 | -82.9242
3.3644 | -82.9356
3.3647 | -82.9469
3.365 | -82.9578
3.3654 | -82.9678
3.3657 | -82.979
3.3661 | -82.9879
3.3664 | -82.9981
3.3667 | -83.0084
3.3671 | -83.0189
3.3674 | -83.0283
3.3678 | -83.0391
3.3681 | -83.0503
3.3684 | -83.0589
3.3688 | -83.0692
3.3691 | -83.081
3.3695 | -83.0871
3.3698 | -83.0986
3.3702 | -83.1096
3.3705 | -83.1193
3.3708 | -83.1291
3.3712 | -83.1382
3.3715 | -83.1475
3.3719 | -83.157
3.3722 | -83.167
3.3725 | -83.1825
3.3729 | -83.1944
3.3732 | -83.2055
3.3736 | -83.2143
3.3739 | -83.2234
3.3742 | -83.2359
3.3746 | -83.2456
3.3749 | -83.2565
3.3753 | -83.2662
3.3756 | -83.2749
3.3759 | -83.2846
3.3763 | -83.2941
3.3766 | -83.307
3.377 | -83.3225
3.3773 | -83.336
3.3776 | -83.3516
3.378 | -83.3691
3.3783 | -83.3815
3.3787 | -83.3898
3.379 | -83.3981
3.3793 | -83.4027
3.3797 | -83.4097
3.38 | -83.4156
3.3804 | -83.4188
3.3807 | -83.427
3.381 | -83.4361
3.3814 | -83.4431
3.3817 | -83.4509
3.3821 | -83.4579
3.3824 | -83.4642
3.3828 | -83.4836
3.3831 | -83.5061
3.3834 | -83.5178
3.3838 | -83.5321
3.3841 | -83.5419
3.3845 | -83.5555
3.3848 | -83.562
3.3851 | -83.5689
3.3855 | -83.5751
3.3858 | -83.5807
3.3862 | -83.5978
3.3865 | -83.6106
3.3868 | -83.6137
3.3872 | -83.6205
3.3875 | -83.6459
3.3879 | -83.6541
3.3882 | -83.6615
3.3885 | -83.6693
3.3889 | -83.6771
3.3892 | -83.6799
3.3896 | -83.6943
3.3899 | -83.7066
3.3902 | -83.7107
3.3906 | -83.7157
3.3909 | -83.7216
3.3913 | -83.7257
3.3916 | -83.7338
3.3919 | -83.7435
3.3923 | -83.7573
3.3926 | -83.7693
3.393 | -83.7743
3.3933 | -83.778
3.3936 | -83.7866
3.394 | -83.7974
3.3943 | -83.8025
3.3947 | -83.8061
3.395 | -83.8096
3.3953 | -83.8198
3.3957 | -83.8268
3.396 | -83.8335
3.3964 | -83.8405
3.3967 | -83.8435
3.3971 | -83.847
3.3974 | -83.85
3.3977 | -83.8525
3.3981 | -83.8553
3.3984 | -83.8578
3.3988 | -83.8574
3.3991 | -83.863
3.3994 | -83.8699
3.3998 | -83.8767
3.4001 | -83.8795
3.4005 | -83.8887
3.4008 | -83.8954
3.4011 | -83.8996
3.4015 | -83.905
3.4018 | -83.9067
3.4022 | -83.907
3.4025 | -83.91
3.4028 | -83.9157
3.4032 | -83.9179
3.4035 | -83.9198
3.4039 | -83.9222
3.4042 | -83.9247
3.4045 | -83.9277
3.4049 | -83.9301
3.4052 | -83.935
3.4056 | -83.9367
3.4059 | -83.9386
3.4062 | -83.9411
3.4066 | -83.9431
3.4069 | -83.9453
3.4073 | -83.9471
3.4076 | -83.95
3.4079 | -83.9522
3.4083 | -83.952
3.4086 | -83.9509
3.409 | -83.9523
3.4093 | -83.9545
3.4096 | -83.9586
3.41 | -83.9613
3.4103 | -83.9634
3.4107 | -83.9651
3.411 | -83.9683
3.4114 | -83.97
3.4117 | -83.972
3.412 | -83.9736
3.4124 | -83.9735
3.4127 | -83.9744
3.4131 | -83.9748
3.4134 | -83.9765
3.4137 | -83.9759
3.4141 | -83.9751
3.4144 | -83.9748
3.4148 | -83.9748
3.4151 | -83.9757
3.4154 | -83.9791
3.4158 | -83.9801
3.4161 | -83.9812
3.4165 | -83.9824
3.4168 | -83.9847
3.4171 | -83.9873
3.4175 | -83.988
3.4178 | -83.9894
3.4182 | -83.9907
3.4185 | -83.9902
3.4188 | -83.9913
3.4192 | -83.9923
3.4195 | -83.9932
3.4199 | -83.9943
3.4202 | -83.9974
3.4205 | -83.9989
3.4209 | -83.999
3.4212 | -84.0001
3.4216 | -84.0021
3.4219 | -84.0037
3.4222 | -84.0056
3.4226 | -84.0096
3.4229 | -84.0125
3.4233 | -84.014
3.4236 | -84.0145
3.4239 | -84.0156
3.4243 | -84.0167
3.4246 | -84.0179
3.425 | -84.0178
3.4253 | -84.0167
3.4257 | -84.0174
3.426 | -84.0168
3.4263 | -84.0194
3.4267 | -84.0201
3.427 | -84.0211
3.4274 | -84.0228
3.4277 | -84.0225
3.428 | -84.0228
3.4284 | -84.0235
3.4287 | -84.0244
3.4291 | -84.0253
3.4294 | -84.0258
3.4297 | -84.0268
3.4301 | -84.0266
3.4304 | -84.0276
3.4308 | -84.0281
3.4311 | -84.0298
3.4314 | -84.0323
3.4318 | -84.0318
3.4321 | -84.0319
3.4325 | -84.0331
3.4328 | -84.0339
3.4331 | -84.0342
3.4335 | -84.0347
3.4338 | -84.0338
3.4342 | -84.0303
3.4345 | -84.029
3.4348 | -84.0291
3.4352 | -84.0301
3.4355 | -84.0317
3.4359 | -84.0322
3.4362 | -84.0322
3.4365 | -84.0327
3.4369 | -84.0331
3.4372 | -84.0339
3.4376 | -84.0341
3.4379 | -84.0344
3.4382 | -84.034
3.4386 | -84.0346
3.4389 | -84.035
3.4393 | -84.036
3.4396 | -84.0365
3.44 | -84.0373
3.4403 | -84.0388
3.4406 | -84.0395
3.441 | -84.038
3.4413 | -84.0385
3.4417 | -84.0381
3.442 | -84.0387
3.4423 | -84.038
3.4427 | -84.0381
3.443 | -84.0384
3.4434 | -84.0386
3.4437 | -84.0367
3.444 | -84.0347
3.4444 | -84.0343
3.4447 | -84.0344
3.4451 | -84.0368
3.4454 | -84.0375
3.4457 | -84.0374
3.4461 | -84.0372
3.4464 | -84.036
3.4468 | -84.0359
3.4471 | -84.0363
3.4474 | -84.0362
3.4478 | -84.0334
3.4481 | -84.0286
3.4485 | -84.0283
3.4488 | -84.0281
3.4491 | -84.0286
3.4495 | -84.0295
3.4498 | -84.0304
3.4502 | -84.03
3.4505 | -84.0291
3.4508 | -84.029
3.4512 | -84.0282
3.4515 | -84.0271
3.4519 | -84.0306
3.4522 | -84.0321
3.4526 | -84.0325
3.4529 | -84.0321
3.4532 | -84.0301
3.4536 | -84.0283
3.4539 | -84.0281
3.4543 | -84.0276
3.4546 | -84.0274
3.4549 | -84.027
3.4553 | -84.0279
3.4556 | -84.0309
3.456 | -84.0318
3.4563 | -84.0303
3.4566 | -84.0277
3.457 | -84.0299
3.4573 | -84.0327
3.4577 | -84.031
3.458 | -84.0313
3.4583 | -84.0304
3.4587 | -84.0299
3.459 | -84.029
3.4594 | -84.0294
3.4597 | -84.0297
3.46 | -84.0291
3.4604 | -84.0291
3.4607 | -84.0291
3.4611 | -84.0284
3.4614 | -84.0242
3.4617 | -84.0216
3.4621 | -84.0214
3.4624 | -84.0205
3.4628 | -84.0215
3.4631 | -84.023
3.4634 | -84.023
3.4638 | -84.0226
3.4641 | -84.0214
3.4645 | -84.0192
3.4648 | -84.0188
3.4651 | -84.0183
3.4655 | -84.0183
3.4658 | -84.0181
3.4662 | -84.0173
3.4665 | -84.016
3.4668 | -84.0167
3.4672 | -84.0184
3.4675 | -84.0187
3.4679 | -84.0176
3.4682 | -84.0171
3.4686 | -84.0179
3.4689 | -84.0187
3.4692 | -84.0186
3.4696 | -84.0182
3.4699 | -84.0182
3.4703 | -84.0184
3.4706 | -84.0184
3.4709 | -84.0183
3.4713 | -84.0174
3.4716 | -84.0171
3.472 | -84.017
3.4723 | -84.017
3.4726 | -84.0168
3.473 | -84.0157
3.4733 | -84.0143
3.4737 | -84.0128
3.474 | -84.0121
3.4743 | -84.0109
3.4747 | -84.0111
3.475 | -84.0095
3.4754 | -84.0125
3.4757 | -84.0145
3.476 | -84.0132
3.4764 | -84.0123
3.4767 | -84.0124
3.4771 | -84.0124
3.4774 | -84.0134
3.4777 | -84.012
3.4781 | -84.0097
3.4784 | -84.0076
3.4788 | -84.0038
3.4791 | -84.0027
3.4794 | -84.0026
3.4798 | -84.0028
3.4801 | -84.0024
3.4805 | -84.0026
3.4808 | -84.0045
3.4812 | -84.005
3.4815 | -84.0052
3.4818 | -84.0053
3.4822 | -84.0043
3.4825 | -84.0028
3.4829 | -84.0023
3.4832 | -84.0014
3.4835 | -84.0007
3.4839 | -84.0001
3.4842 | -83.9991
3.4846 | -83.9974
3.4849 | -83.9972
3.4852 | -83.9981
3.4856 | -83.9975
3.4859 | -83.9998
3.4863 | -84.0029
3.4866 | -84.0026
3.4869 | -83.9994
3.4873 | -83.9985
3.4876 | -83.9984
3.488 | -83.9977
3.4883 | -83.9975
3.4886 | -83.9943
3.489 | -83.9932
3.4893 | -83.9927
3.4897 | -83.9915
3.49 | -83.9902
3.4903 | -83.9901
3.4907 | -83.9899
3.491 | -83.9898
3.4914 | -83.9902
3.4917 | -83.9925
3.492 | -83.991
3.4924 | -83.994
3.4927 | -83.9959
3.4931 | -83.9955
3.4934 | -83.9949
3.4937 | -83.9945
3.4941 | -83.9938
3.4944 | -83.9917
3.4948 | -83.9898
3.4951 | -83.9881
3.4955 | -83.9868
3.4958 | -83.9856
3.4961 | -83.984
3.4965 | -83.9829
3.4968 | -83.9817
3.4972 | -83.9823
3.4975 | -83.9836
3.4978 | -83.9869
3.4982 | -83.9873
3.4985 | -83.9863
3.4989 | -83.985
3.4992 | -83.9833
3.4995 | -83.9828
3.4999 | -83.9818
3.5002 | -83.982
3.5006 | -83.9837
3.5009 | -83.9839
3.5012 | -83.9837
3.5016 | -83.9836
3.5019 | -83.9828
3.5023 | -83.9807
3.5026 | -83.9787
3.5029 | -83.9778
3.5033 | -83.9755
3.5036 | -83.9737
3.504 | -83.9735
3.5043 | -83.9741
3.5046 | -83.974
3.505 | -83.9723
3.5053 | -83.9697
3.5057 | -83.9667
3.506 | -83.9642
3.5063 | -83.9631
3.5067 | -83.9618
3.507 | -83.962
3.5074 | -83.9606
3.5077 | -83.9589
3.508 | -83.9585
3.5084 | -83.9583
3.5087 | -83.9576
3.5091 | -83.9576
3.5094 | -83.9586
3.5098 | -83.9587
3.5101 | -83.9579
3.5104 | -83.9579
3.5108 | -83.9572
3.5111 | -83.9546
3.5115 | -83.9497
3.5118 | -83.9465
3.5121 | -83.9467
3.5125 | -83.9456
3.5128 | -83.9438
3.5132 | -83.9424
3.5135 | -83.9415
3.5138 | -83.9413
3.5142 | -83.9414
3.5145 | -83.9407
3.5149 | -83.9402
3.5152 | -83.9392
3.5155 | -83.9371
3.5159 | -83.9355
3.5162 | -83.9362
3.5166 | -83.9369
3.5169 | -83.9368
3.5172 | -83.9361
3.5176 | -83.9361
3.5179 | -83.9368
3.5183 | -83.9363
3.5186 | -83.9343
3.5189 | -83.9316
3.5193 | -83.93
3.5196 | -83.9287
3.52 | -83.9276
3.5203 | -83.9257
3.5206 | -83.9251
3.521 | -83.9244
3.5213 | -83.9236
3.5217 | -83.9218
3.522 | -83.9203
3.5223 | -83.9203
3.5227 | -83.9206
3.523 | -83.9199
3.5234 | -83.9194
3.5237 | -83.9189
3.5241 | -83.918
3.5244 | -83.9173
3.5247 | -83.9159
3.5251 | -83.915
3.5254 | -83.9149
3.5258 | -83.9146
3.5261 | -83.914
3.5264 | -83.9123
3.5268 | -83.9117
3.5271 | -83.9095
3.5275 | -83.9063
3.5278 | -83.9057
3.5281 | -83.9054
3.5285 | -83.9069
3.5288 | -83.9051
3.5292 | -83.9041
3.5295 | -83.9032
3.5298 | -83.9021
3.5302 | -83.8998
3.5305 | -83.897
3.5309 | -83.8964
3.5312 | -83.8973
3.5315 | -83.9001
3.5319 | -83.9001
3.5322 | -83.8987
3.5326 | -83.8974
3.5329 | -83.8968
3.5332 | -83.8956
3.5336 | -83.8939
3.5339 | -83.8939
3.5343 | -83.891
3.5346 | -83.8883
3.5349 | -83.8868
3.5353 | -83.883
3.5356 | -83.8822
3.536 | -83.881
3.5363 | -83.8795
3.5366 | -83.8776
3.537 | -83.8762
3.5373 | -83.8748
3.5377 | -83.8725
3.538 | -83.871
3.5384 | -83.8698
3.5387 | -83.8683
3.539 | -83.867
3.5394 | -83.866
3.5397 | -83.8646
3.5401 | -83.864
3.5404 | -83.8628
3.5407 | -83.8614
3.5411 | -83.8602
3.5414 | -83.8593
3.5418 | -83.8571
3.5421 | -83.8544
3.5424 | -83.8535
3.5428 | -83.852
3.5431 | -83.8498
3.5435 | -83.8486
3.5438 | -83.8472
3.5441 | -83.8462
3.5445 | -83.8449
3.5448 | -83.8421
3.5452 | -83.8401
3.5455 | -83.8387
3.5458 | -83.8364
3.5462 | -83.8325
3.5465 | -83.8322
3.5469 | -83.8314
3.5472 | -83.8295
3.5475 | -83.827
3.5479 | -83.8254
3.5482 | -83.8233
3.5486 | -83.8214
3.5489 | -83.8216
3.5492 | -83.8204
3.5496 | -83.8203
3.5499 | -83.8185
3.5503 | -83.8159
3.5506 | -83.8141
3.551 | -83.8115
3.5513 | -83.8092
3.5516 | -83.8075
3.552 | -83.8081
3.5523 | -83.8082
3.5527 | -83.8069
3.553 | -83.804
3.5533 | -83.8013
3.5537 | -83.7992
3.554 | -83.7976
3.5544 | -83.796
3.5547 | -83.7949
3.555 | -83.7946
3.5554 | -83.794
3.5557 | -83.7926
3.5561 | -83.792
3.5564 | -83.7906
3.5567 | -83.7891
3.5571 | -83.7867
3.5574 | -83.7827
3.5578 | -83.7802
3.5581 | -83.7782
3.5584 | -83.7763
3.5588 | -83.7746
3.5591 | -83.7722
3.5595 | -83.7693
3.5598 | -83.7674
3.5601 | -83.7661
3.5605 | -83.7648
3.5608 | -83.7633
3.5612 | -83.7619
3.5615 | -83.7601
3.5618 | -83.7584
3.5622 | -83.7582
3.5625 | -83.7571
3.5629 | -83.7561
3.5632 | -83.7556
3.5635 | -83.7545
3.5639 | -83.7536
3.5642 | -83.7513
3.5646 | -83.7515
3.5649 | -83.7517
3.5653 | -83.7501
3.5656 | -83.7482
3.5659 | -83.747
3.5663 | -83.7444
3.5666 | -83.7423
3.567 | -83.7404
3.5673 | -83.7389
3.5676 | -83.7378
3.568 | -83.7363
3.5683 | -83.736
3.5687 | -83.7352
3.569 | -83.734
3.5693 | -83.7324
3.5697 | -83.7321
3.57 | -83.7328
3.5704 | -83.7313
3.5707 | -83.73
3.571 | -83.7292
3.5714 | -83.7272
3.5717 | -83.7257
3.5721 | -83.7252
3.5724 | -83.7237
3.5727 | -83.7218
3.5731 | -83.7197
3.5734 | -83.7186
3.5738 | -83.7174
3.5741 | -83.7162
3.5744 | -83.715
3.5748 | -83.7134
3.5751 | -83.712
3.5755 | -83.7107
3.5758 | -83.7101
3.5761 | -83.7092
3.5765 | -83.7081
3.5768 | -83.7073
3.5772 | -83.7061
3.5775 | -83.7055
3.5778 | -83.7045
3.5782 | -83.7018
3.5785 | -83.7004
3.5789 | -83.6988
3.5792 | -83.6981
3.5796 | -83.6958
3.5799 | -83.693
3.5802 | -83.6918
3.5806 | -83.6906
3.5809 | -83.6893
3.5813 | -83.6877
3.5816 | -83.6857
3.5819 | -83.6847
3.5823 | -83.683
3.5826 | -83.6814
3.583 | -83.6802
3.5833 | -83.6807
3.5836 | -83.6804
3.584 | -83.6792
3.5843 | -83.678
3.5847 | -83.6769
3.585 | -83.6764
3.5853 | -83.6753
3.5857 | -83.676
3.586 | -83.6752
3.5864 | -83.6742
3.5867 | -83.6725
3.587 | -83.6709
3.5874 | -83.6702
3.5877 | -83.6713
3.5881 | -83.6732
3.5884 | -83.6727
3.5887 | -83.6725
3.5891 | -83.6713
3.5894 | -83.6708
3.5898 | -83.6703
3.5901 | -83.6701
3.5904 | -83.6708
3.5908 | -83.6703
3.5911 | -83.6693
3.5915 | -83.6685
3.5918 | -83.6686
3.5922 | -83.6693
3.5925 | -83.6679
3.5928 | -83.6647
3.5932 | -83.6641
3.5935 | -83.665
3.5939 | -83.6642
3.5942 | -83.6622
3.5945 | -83.6602
3.5949 | -83.659
3.5952 | -83.6605
3.5956 | -83.6603
3.5959 | -83.6601
3.5962 | -83.6599
3.5966 | -83.6599
3.5969 | -83.6594
3.5973 | -83.6586
3.5976 | -83.6585
3.5979 | -83.6585
3.5983 | -83.6579
3.5986 | -83.6574
3.599 | -83.6566
3.5993 | -83.6552
3.5996 | -83.6549
3.6 | -83.6548
3.6003 | -83.6568
3.6007 | -83.6582
3.601 | -83.6577
3.6013 | -83.6556
3.6017 | -83.6537
3.602 | -83.6525
3.6024 | -83.6526
3.6027 | -83.6525
3.603 | -83.6519
3.6034 | -83.651
3.6037 | -83.65
3.6041 | -83.649
3.6044 | -83.6486
3.6047 | -83.6488
3.6051 | -83.6508
3.6054 | -83.6505
3.6058 | -83.6508
3.6061 | -83.6519
3.6065 | -83.6515
3.6068 | -83.6505
3.6071 | -83.6493
3.6075 | -83.6485
3.6078 | -83.6489
3.6082 | -83.6496
3.6085 | -83.6517
3.6088 | -83.6533
3.6092 | -83.6535
3.6095 | -83.6537
3.6099 | -83.6546
3.6102 | -83.6536
3.6105 | -83.6532
3.6109 | -83.6552
3.6112 | -83.6563
3.6116 | -83.6562
3.6119 | -83.6561
3.6122 | -83.6563
3.6126 | -83.6576
3.6129 | -83.6583
3.6133 | -83.6596
3.6136 | -83.6599
3.6139 | -83.6606
3.6143 | -83.6607
3.6146 | -83.6614
3.615 | -83.6643
3.6153 | -83.6648
3.6156 | -83.6654
3.616 | -83.6658
3.6163 | -83.6663
3.6167 | -83.667
3.617 | -83.6672
3.6173 | -83.6678
3.6177 | -83.6678
3.618 | -83.6677
3.6184 | -83.6684
3.6187 | -83.6683
3.619 | -83.6691
3.6194 | -83.67
3.6197 | -83.6708
3.6201 | -83.6717
3.6204 | -83.6728
3.6208 | -83.674
3.6211 | -83.6738
3.6214 | -83.6733
3.6218 | -83.6738
3.6221 | -83.6753
3.6225 | -83.6765
3.6228 | -83.6776
3.6231 | -83.6802
3.6235 | -83.6849
3.6238 | -83.6899
3.6242 | -83.6919
3.6245 | -83.6953
3.6248 | -83.6974
3.6252 | -83.7011
3.6255 | -83.7034
3.6259 | -83.705
3.6262 | -83.7063
3.6265 | -83.7068
3.6269 | -83.7078
3.6272 | -83.7088
3.6276 | -83.7104
3.6279 | -83.7137
3.6282 | -83.7158
3.6286 | -83.7171
3.6289 | -83.7182
3.6293 | -83.7196
3.6296 | -83.7208
3.6299 | -83.7218
3.6303 | -83.7236
3.6306 | -83.7275
3.631 | -83.7311
3.6313 | -83.7329
3.6316 | -83.7344
3.632 | -83.7364
3.6323 | -83.7381
3.6327 | -83.7401
3.633 | -83.7396
3.6333 | -83.7398
3.6337 | -83.7416
3.634 | -83.7424
3.6344 | -83.7442
3.6347 | -83.7476
3.6351 | -83.7503
3.6354 | -83.7525
3.6357 | -83.7542
3.6361 | -83.7561
3.6364 | -83.7576
3.6368 | -83.7579
3.6371 | -83.7587
3.6374 | -83.7599
3.6378 | -83.7616
3.6381 | -83.7633
3.6385 | -83.7646
3.6388 | -83.7664
3.6391 | -83.7681
3.6395 | -83.7692
3.6398 | -83.7699
3.6402 | -83.7715
3.6405 | -83.7716
3.6408 | -83.7719
3.6412 | -83.7734
3.6415 | -83.7741
3.6419 | -83.7755
3.6422 | -83.7777
3.6425 | -83.78
3.6429 | -83.783
3.6432 | -83.787
3.6436 | -83.7897
3.6439 | -83.7917
3.6442 | -83.7938
3.6446 | -83.7979
3.6449 | -83.7997
3.6453 | -83.8001
3.6456 | -83.8024
3.6459 | -83.8048
3.6463 | -83.8071
3.6466 | -83.8086
3.647 | -83.8113
3.6473 | -83.8158
3.6476 | -83.8183
3.648 | -83.8208
3.6483 | -83.8233
3.6487 | -83.8258
3.649 | -83.8279
3.6494 | -83.8298
3.6497 | -83.8325
3.65 | -83.835
3.6504 | -83.8377
3.6507 | -83.8403
3.6511 | -83.843
3.6514 | -83.8458
3.6517 | -83.8489
3.6521 | -83.8522
3.6524 | -83.8546
3.6528 | -83.8576
3.6531 | -83.8612
3.6534 | -83.8655
3.6538 | -83.8697
3.6541 | -83.8729
3.6545 | -83.8777
3.6548 | -83.8806
3.6551 | -83.8843
3.6555 | -83.888
3.6558 | -83.8912
3.6562 | -83.8941
3.6565 | -83.8976
3.6568 | -83.9004
3.6572 | -83.9036
3.6575 | -83.9063
3.6579 | -83.9083
3.6582 | -83.9113
3.6585 | -83.9123
3.6589 | -83.9144
3.6592 | -83.9164
3.6596 | -83.9198
3.6599 | -83.9239
3.6602 | -83.9261
3.6606 | -83.9281
3.6609 | -83.9309
3.6613 | -83.9318
3.6616 | -83.9323
3.662 | -83.9349
3.6623 | -83.9378
3.6626 | -83.9412
3.663 | -83.9442
3.6633 | -83.947
3.6637 | -83.9498
3.664 | -83.9529
3.6643 | -83.9555
3.6647 | -83.9591
3.665 | -83.963
3.6654 | -83.9654
3.6657 | -83.967
3.666 | -83.9683
3.6664 | -83.971
3.6667 | -83.9723
3.6671 | -83.975
3.6674 | -83.9785
3.6677 | -83.9815
3.6681 | -83.9844
3.6684 | -83.9871
3.6688 | -83.9899
3.6691 | -83.9928
3.6694 | -83.9958
3.6698 | -83.9989
3.6701 | -84.002
3.6705 | -84.0048
3.6708 | -84.0073
3.6711 | -84.007
3.6715 | -84.0085
3.6718 | -84.0138
3.6722 | -84.0177
3.6725 | -84.0226
3.6728 | -84.0244
3.6732 | -84.0253
3.6735 | -84.0273
3.6739 | -84.0301
3.6742 | -84.034
3.6745 | -84.038
3.6749 | -84.041
3.6752 | -84.0437
3.6756 | -84.0436
3.6759 | -84.0444
3.6763 | -84.047
3.6766 | -84.0501
3.6769 | -84.0524
3.6773 | -84.0552
3.6776 | -84.058
3.678 | -84.0615
3.6783 | -84.0643
3.6786 | -84.0671
3.679 | -84.0716
3.6793 | -84.0744
3.6797 | -84.0774
3.68 | -84.0804
3.6803 | -84.084
3.6807 | -84.0877
3.681 | -84.0926
3.6814 | -84.0993
3.6817 | -84.1046
3.682 | -84.1067
3.6824 | -84.1085
3.6827 | -84.1115
3.6831 | -84.1136
3.6834 | -84.1152
3.6837 | -84.1157
3.6841 | -84.1178
3.6844 | -84.12
3.6848 | -84.1228
3.6851 | -84.1254
3.6854 | -84.1276
3.6858 | -84.1306
3.6861 | -84.134
3.6865 | -84.1358
3.6868 | -84.137
3.6871 | -84.1386
3.6875 | -84.1404
3.6878 | -84.143
3.6882 | -84.1453
3.6885 | -84.1492
3.6888 | -84.1525
3.6892 | -84.1544
3.6895 | -84.1574
3.6899 | -84.1601
3.6902 | -84.1635
3.6906 | -84.166
3.6909 | -84.1692
3.6912 | -84.1723
3.6916 | -84.1769
3.6919 | -84.1813
3.6923 | -84.1835
3.6926 | -84.1861
3.6929 | -84.1887
3.6933 | -84.1909
3.6936 | -84.1933
3.694 | -84.1959
3.6943 | -84.1984
3.6946 | -84.2014
3.695 | -84.2044
3.6953 | -84.2061
3.6957 | -84.208
3.696 | -84.2124
3.6963 | -84.2158
3.6967 | -84.2171
3.697 | -84.2185
3.6974 | -84.2212
3.6977 | -84.2231
3.698 | -84.2239
3.6984 | -84.2252
3.6987 | -84.2274
3.6991 | -84.2301
3.6994 | -84.2319
3.6997 | -84.2345
3.7001 | -84.2377
3.7004 | -84.2395
3.7008 | -84.2422
3.7011 | -84.2444
3.7014 | -84.2457
3.7018 | -84.2472
3.7021 | -84.2487
3.7025 | -84.2503
3.7028 | -84.252
3.7031 | -84.2531
3.7035 | -84.2541
3.7038 | -84.2549
3.7042 | -84.2559
3.7045 | -84.2573
3.7049 | -84.2586
3.7052 | -84.2601
3.7055 | -84.2642
3.7059 | -84.2673
3.7062 | -84.2674
3.7066 | -84.2678
3.7069 | -84.2683
3.7072 | -84.2696
3.7076 | -84.2711
3.7079 | -84.273
3.7083 | -84.2754
3.7086 | -84.2768
3.7089 | -84.2779
3.7093 | -84.2784
3.7096 | -84.2784
3.71 | -84.2786
3.7103 | -84.2808
3.7106 | -84.2854
3.711 | -84.2884
3.7113 | -84.2893
3.7117 | -84.291
3.712 | -84.292
3.7123 | -84.2934
3.7127 | -84.294
3.713 | -84.2943
3.7134 | -84.2945
3.7137 | -84.2964
3.714 | -84.2984
3.7144 | -84.2997
3.7147 | -84.301
3.7151 | -84.3016
3.7154 | -84.3026
3.7157 | -84.3043
3.7161 | -84.3061
3.7164 | -84.3067
3.7168 | -84.3069
3.7171 | -84.3063
3.7174 | -84.3054
3.7178 | -84.306
3.7181 | -84.3067
3.7185 | -84.3066
3.7188 | -84.3068
3.7192 | -84.3066
3.7195 | -84.3076
3.7198 | -84.3071
3.7202 | -84.3068
3.7205 | -84.3074
3.7209 | -84.3094
3.7212 | -84.3111
3.7215 | -84.3105
3.7219 | -84.3103
3.7222 | -84.3114
3.7226 | -84.3118
3.7229 | -84.313
3.7232 | -84.3132
3.7236 | -84.3146
3.7239 | -84.315
3.7243 | -84.3134
3.7246 | -84.3127
3.7249 | -84.3133
3.7253 | -84.3148
3.7256 | -84.3153
3.726 | -84.3154
3.7263 | -84.3173
3.7266 | -84.3179
3.727 | -84.3182
3.7273 | -84.3183
3.7277 | -84.3187
3.728 | -84.3186
3.7283 | -84.3203
3.7287 | -84.3232
3.729 | -84.3257
3.7294 | -84.3286
3.7297 | -84.328
3.73 | -84.3262
3.7304 | -84.3248
3.7307 | -84.3239
3.7311 | -84.3234
3.7314 | -84.323
3.7317 | -84.3231
3.7321 | -84.3239
3.7324 | -84.3233
3.7328 | -84.3204
3.7331 | -84.3211
3.7335 | -84.3218
3.7338 | -84.3221
3.7341 | -84.3221
3.7345 | -84.3209
3.7348 | -84.3205
3.7352 | -84.3204
3.7355 | -84.3213
3.7358 | -84.3211
3.7362 | -84.3207
3.7365 | -84.3207
3.7369 | -84.3202
3.7372 | -84.3197
3.7375 | -84.319
3.7379 | -84.3195
3.7382 | -84.3196
3.7386 | -84.319
3.7389 | -84.3179
3.7392 | -84.3176
3.7396 | -84.3171
3.7399 | -84.3192
3.7403 | -84.3197
3.7406 | -84.3191
3.7409 | -84.3189
3.7413 | -84.3184
3.7416 | -84.3174
3.742 | -84.3167
3.7423 | -84.3161
3.7426 | -84.3152
3.743 | -84.3151
3.7433 | -84.3149
3.7437 | -84.3162
3.744 | -84.3163
3.7443 | -84.3159
3.7447 | -84.3156
3.745 | -84.3151
3.7454 | -84.3144
3.7457 | -84.3144
3.746 | -84.316
3.7464 | -84.3172
3.7467 | -84.3152
3.7471 | -84.3135
3.7474 | -84.3124
3.7478 | -84.3113
3.7481 | -84.3112
3.7484 | -84.3113
3.7488 | -84.3106
3.7491 | -84.3094
3.7495 | -84.3075
3.7498 | -84.3066
3.7501 | -84.3059
3.7505 | -84.3053
3.7508 | -84.3048
3.7512 | -84.3047
3.7515 | -84.3034
3.7518 | -84.3017
3.7522 | -84.2992
3.7525 | -84.2976
3.7529 | -84.2947
3.7532 | -84.2919
3.7535 | -84.2904
3.7539 | -84.2891
3.7542 | -84.2875
3.7546 | -84.288
3.7549 | -84.2876
3.7552 | -84.2875
3.7556 | -84.2881
3.7559 | -84.2878
3.7563 | -84.2862
3.7566 | -84.2847
3.7569 | -84.283
3.7573 | -84.2819
3.7576 | -84.2816
3.758 | -84.2838
3.7583 | -84.2842
3.7586 | -84.2828
3.759 | -84.2812
3.7593 | -84.2798
3.7597 | -84.2784
3.76 | -84.2775
3.7604 | -84.2767
3.7607 | -84.2752
3.761 | -84.2738
3.7614 | -84.2725
3.7617 | -84.2702
3.7621 | -84.2676
3.7624 | -84.2662
3.7627 | -84.2652
3.7631 | -84.264
3.7634 | -84.2621
3.7638 | -84.2609
3.7641 | -84.2618
3.7644 | -84.2647
3.7648 | -84.2659
3.7651 | -84.2635
3.7655 | -84.2619
3.7658 | -84.2609
3.7661 | -84.2595
3.7665 | -84.2579
3.7668 | -84.2572
3.7672 | -84.2558
3.7675 | -84.2544
3.7678 | -84.2526
3.7682 | -84.2513
3.7685 | -84.2507
3.7689 | -84.2505
3.7692 | -84.2501
3.7695 | -84.248
3.7699 | -84.246
3.7702 | -84.2448
3.7706 | -84.2432
3.7709 | -84.2416
3.7712 | -84.2388
3.7716 | -84.2348
3.7719 | -84.232
3.7723 | -84.2281
3.7726 | -84.2292
3.7729 | -84.2289
3.7733 | -84.2266
3.7736 | -84.225
3.774 | -84.2245
3.7743 | -84.2216
3.7747 | -84.2201
3.775 | -84.2179
3.7753 | -84.2176
3.7757 | -84.2176
3.776 | -84.2162
3.7764 | -84.2148
3.7767 | -84.2134
3.777 | -84.212
3.7774 | -84.2107
3.7777 | -84.2093
3.7781 | -84.207
3.7784 | -84.2058
3.7787 | -84.2038
3.7791 | -84.2021
3.7794 | -84.2014
3.7798 | -84.1998
3.7801 | -84.1992
3.7804 | -84.1979
3.7808 | -84.1961
3.7811 | -84.1959
3.7815 | -84.1967
3.7818 | -84.1952
3.7821 | -84.1951
3.7825 | -84.193
3.7828 | -84.1914
3.7832 | -84.1914
3.7835 | -84.19
3.7838 | -84.1879
3.7842 | -84.186
3.7845 | -84.1853
3.7849 | -84.1846
3.7852 | -84.1814
3.7855 | -84.1786
3.7859 | -84.1764
3.7862 | -84.1749
3.7866 | -84.1741
3.7869 | -84.173
3.7872 | -84.1718
3.7876 | -84.1707
3.7879 | -84.1707
3.7883 | -84.1686
3.7886 | -84.1693
3.789 | -84.1694
3.7893 | -84.1679
3.7896 | -84.1687
3.79 | -84.1676
3.7903 | -84.1665
3.7907 | -84.1655
3.791 | -84.1642
3.7913 | -84.1632
3.7917 | -84.1623
3.792 | -84.1624
3.7924 | -84.1619
3.7927 | -84.1622
3.793 | -84.1621
3.7934 | -84.1603
3.7937 | -84.159
3.7941 | -84.1574
3.7944 | -84.1558
3.7947 | -84.1532
3.7951 | -84.1515
3.7954 | -84.1512
3.7958 | -84.1495
3.7961 | -84.1464
3.7964 | -84.1447
3.7968 | -84.1444
3.7971 | -84.1439
3.7975 | -84.1417
3.7978 | -84.1401
3.7981 | -84.1401
3.7985 | -84.1397
3.7988 | -84.1387
3.7992 | -84.1382
3.7995 | -84.1369
3.7998 | -84.1377
3.8002 | -84.1365
3.8005 | -84.1353
3.8009 | -84.1355
3.8012 | -84.136
3.8016 | -84.1329
3.8019 | -84.131
3.8022 | -84.1289
3.8026 | -84.1273
3.8029 | -84.1267
3.8033 | -84.1263
3.8036 | -84.1248
3.8039 | -84.1246
3.8043 | -84.1252
3.8046 | -84.1238
3.805 | -84.1227
3.8053 | -84.1214
3.8056 | -84.1209
3.806 | -84.1206
3.8063 | -84.1201
3.8067 | -84.1199
3.807 | -84.1195
3.8073 | -84.1169
3.8077 | -84.1145
3.808 | -84.1133
3.8084 | -84.1126
3.8087 | -84.1129
3.809 | -84.112
3.8094 | -84.1117
3.8097 | -84.1109
3.8101 | -84.1097
3.8104 | -84.1087
3.8107 | -84.1066
3.8111 | -84.1053
3.8114 | -84.1047
3.8118 | -84.1021
3.8121 | -84.1009
3.8124 | -84.0999
3.8128 | -84.0978
3.8131 | -84.0965
3.8135 | -84.0968
3.8138 | -84.0986
3.8141 | -84.0994
3.8145 | -84.0993
3.8148 | -84.0989
3.8152 | -84.0986
3.8155 | -84.0983
3.8159 | -84.0986
3.8162 | -84.0981
3.8165 | -84.0967
3.8169 | -84.096
3.8172 | -84.0945
3.8176 | -84.0918
3.8179 | -84.0906
3.8182 | -84.0892
3.8186 | -84.0886
3.8189 | -84.0889
3.8193 | -84.0896
3.8196 | -84.0891
3.8199 | -84.0887
3.8203 | -84.0883
3.8206 | -84.0884
3.821 | -84.0881
3.8213 | -84.0885
3.8216 | -84.0877
3.822 | -84.088
3.8223 | -84.0874
3.8227 | -84.0868
3.823 | -84.0868
3.8233 | -84.0868
3.8237 | -84.0845
3.824 | -84.083
3.8244 | -84.0823
3.8247 | -84.0817
3.825 | -84.0806
3.8254 | -84.0797
3.8257 | -84.0783
3.8261 | -84.0777
3.8264 | -84.0777
3.8267 | -84.0778
3.8271 | -84.0764
3.8274 | -84.0758
3.8278 | -84.076
3.8281 | -84.0753
3.8284 | -84.0741
3.8288 | -84.0734
3.8291 | -84.0732
3.8295 | -84.0733
3.8298 | -84.0735
3.8302 | -84.0736
3.8305 | -84.0729
3.8308 | -84.0717
3.8312 | -84.0714
3.8315 | -84.0706
3.8319 | -84.0704
3.8322 | -84.0702
3.8325 | -84.0682
3.8329 | -84.0659
3.8332 | -84.064
3.8336 | -84.0638
3.8339 | -84.0638
3.8342 | -84.0633
3.8346 | -84.063
3.8349 | -84.0628
3.8353 | -84.0629
3.8356 | -84.0633
3.8359 | -84.0645
3.8363 | -84.0644
3.8366 | -84.0641
3.837 | -84.0632
3.8373 | -84.0616
3.8376 | -84.0611
3.838 | -84.0599
3.8383 | -84.0581
3.8387 | -84.0575
3.839 | -84.0585
3.8393 | -84.0595
3.8397 | -84.0587
3.84 | -84.0591
3.8404 | -84.0592
3.8407 | -84.059
3.841 | -84.0601
3.8414 | -84.06
3.8417 | -84.0594
3.8421 | -84.0594
3.8424 | -84.0594
3.8427 | -84.0593
3.8431 | -84.0592
3.8434 | -84.059
3.8438 | -84.0581
3.8441 | -84.0578
3.8445 | -84.0577
3.8448 | -84.0573
3.8451 | -84.0576
3.8455 | -84.0575
3.8458 | -84.0575
3.8462 | -84.0565
3.8465 | -84.0548
3.8468 | -84.0527
3.8472 | -84.0502
3.8475 | -84.0506
3.8479 | -84.0487
3.8482 | -84.0466
3.8485 | -84.0462
3.8489 | -84.0464
3.8492 | -84.0463
3.8496 | -84.0452
3.8499 | -84.0437
3.8502 | -84.0432
3.8506 | -84.0429
3.8509 | -84.0432
3.8513 | -84.0432
3.8516 | -84.0445
3.8519 | -84.0445
3.8523 | -84.0439
3.8526 | -84.0444
3.853 | -84.0459
3.8533 | -84.0469
3.8536 | -84.0471
3.854 | -84.0465
3.8543 | -84.0453
3.8547 | -84.045
3.855 | -84.0448
3.8553 | -84.0449
3.8557 | -84.0448
3.856 | -84.0447
3.8564 | -84.0443
3.8567 | -84.0423
3.857 | -84.0408
3.8574 | -84.0408
3.8577 | -84.0411
3.8581 | -84.0419
3.8584 | -84.0427
3.8588 | -84.0421
3.8591 | -84.0422
3.8594 | -84.0425
3.8598 | -84.0425
3.8601 | -84.0429
3.8605 | -84.0453
3.8608 | -84.0462
3.8611 | -84.0483
3.8615 | -84.049
3.8618 | -84.0492
3.8622 | -84.0497
3.8625 | -84.0494
3.8628 | -84.0497
3.8632 | -84.0502
3.8635 | -84.0502
3.8639 | -84.0497
3.8642 | -84.0489
3.8645 | -84.0488
3.8649 | -84.0479
3.8652 | -84.0467
3.8656 | -84.0458
3.8659 | -84.0458
3.8662 | -84.0462
3.8666 | -84.046
3.8669 | -84.0466
3.8673 | -84.0476
3.8676 | -84.0475
3.8679 | -84.0465
3.8683 | -84.0453
3.8686 | -84.0431
3.869 | -84.042
3.8693 | -84.042
3.8696 | -84.0421
3.87 | -84.0429
3.8703 | -84.043
3.8707 | -84.0447
3.871 | -84.0453
3.8714 | -84.0454
3.8717 | -84.0449
3.872 | -84.043
3.8724 | -84.0408
3.8727 | -84.0397
3.8731 | -84.0395
3.8734 | -84.0397
3.8737 | -84.0388
3.8741 | -84.0388
3.8744 | -84.0391
3.8748 | -84.0386
3.8751 | -84.0385
3.8754 | -84.0383
3.8758 | -84.0378
3.8761 | -84.0378
3.8765 | -84.0381
3.8768 | -84.0383
3.8771 | -84.0381
3.8775 | -84.0384
3.8778 | -84.0376
3.8782 | -84.0381
3.8785 | -84.0389
3.8788 | -84.0393
3.8792 | -84.0401
3.8795 | -84.0401
3.8799 | -84.0395
3.8802 | -84.0397
3.8805 | -84.0386
3.8809 | -84.0377
3.8812 | -84.0376
3.8816 | -84.0375
3.8819 | -84.0379
3.8822 | -84.0386
3.8826 | -84.0386
3.8829 | -84.0378
3.8833 | -84.0371
3.8836 | -84.0366
3.8839 | -84.0367
3.8843 | -84.0365
3.8846 | -84.0368
3.885 | -84.0362
3.8853 | -84.0353
3.8857 | -84.0357
3.886 | -84.0372
3.8863 | -84.0376
3.8867 | -84.0376
3.887 | -84.0374
3.8874 | -84.0376
3.8877 | -84.0388
3.888 | -84.0377
3.8884 | -84.0353
3.8887 | -84.0356
3.8891 | -84.0355
3.8894 | -84.0356
3.8897 | -84.0362
3.8901 | -84.0364
3.8904 | -84.0354
3.8908 | -84.0355
3.8911 | -84.0357
3.8914 | -84.0362
3.8918 | -84.0364
3.8921 | -84.0363
3.8925 | -84.0369
3.8928 | -84.038
3.8931 | -84.0382
3.8935 | -84.0382
3.8938 | -84.0384
3.8942 | -84.0387
3.8945 | -84.0411
3.8948 | -84.0428
3.8952 | -84.043
3.8955 | -84.044
3.8959 | -84.0443
3.8962 | -84.0451
3.8965 | -84.0444
3.8969 | -84.0435
3.8972 | -84.0431
3.8976 | -84.042
3.8979 | -84.0419
3.8982 | -84.0427
3.8986 | -84.0442
3.8989 | -84.0445
3.8993 | -84.0437
3.8996 | -84.043
3.9 | -84.0441
3.9003 | -84.0447
3.9006 | -84.0458
3.901 | -84.0458
3.9013 | -84.0459
3.9017 | -84.0476
3.902 | -84.0493
3.9023 | -84.0497
3.9027 | -84.0518
3.903 | -84.0532
3.9034 | -84.0541
3.9037 | -84.0556
3.904 | -84.0566
3.9044 | -84.0562
3.9047 | -84.0562
3.9051 | -84.0562
3.9054 | -84.0559
3.9057 | -84.0559
3.9061 | -84.0564
3.9064 | -84.0565
3.9068 | -84.0565
3.9071 | -84.0567
3.9074 | -84.0575
3.9078 | -84.0595
3.9081 | -84.0602
3.9085 | -84.0595
3.9088 | -84.0593
3.9091 | -84.0598
3.9095 | -84.0602
3.9098 | -84.0619
3.9102 | -84.0616
3.9105 | -84.0621
3.9108 | -84.0616
3.9112 | -84.0615
3.9115 | -84.0616
3.9119 | -84.0621
3.9122 | -84.0626
3.9125 | -84.0626
3.9129 | -84.063
3.9132 | -84.0631
3.9136 | -84.0625
3.9139 | -84.0629
3.9143 | -84.0626
3.9146 | -84.0617
3.9149 | -84.0624
3.9153 | -84.0631
3.9156 | -84.0641
3.916 | -84.0645
3.9163 | -84.0651
3.9166 | -84.0662
3.917 | -84.0667
3.9173 | -84.0666
3.9177 | -84.0668
3.918 | -84.0673
3.9183 | -84.0652
3.9187 | -84.0627
3.919 | -84.0629
3.9194 | -84.0638
3.9197 | -84.0639
3.92 | -84.0646
3.9204 | -84.0654
3.9207 | -84.0659
3.9211 | -84.0668
3.9214 | -84.0685
3.9217 | -84.0698
3.9221 | -84.0716
3.9224 | -84.0716
3.9228 | -84.0724
3.9231 | -84.0741
3.9234 | -84.0747
3.9238 | -84.0759
3.9241 | -84.075
3.9245 | -84.0717
3.9248 | -84.0726
3.9251 | -84.0756
3.9255 | -84.0763
3.9258 | -84.0765
3.9262 | -84.0767
3.9265 | -84.0761
3.9268 | -84.0758
3.9272 | -84.0757
3.9275 | -84.0766
3.9279 | -84.0774
3.9282 | -84.0773
3.9286 | -84.0766
3.9289 | -84.077
3.9292 | -84.0766
3.9296 | -84.0774
3.9299 | -84.079
3.9303 | -84.0807
3.9306 | -84.0814
3.9309 | -84.0825
3.9313 | -84.0826
3.9316 | -84.0831
3.932 | -84.0838
3.9323 | -84.0844
3.9326 | -84.085
3.933 | -84.0852
3.9333 | -84.0855
3.9337 | -84.0848
3.934 | -84.0848
3.9343 | -84.0853
3.9347 | -84.0858
3.935 | -84.0871
3.9354 | -84.0902
3.9357 | -84.0919
3.936 | -84.0915
3.9364 | -84.092
3.9367 | -84.091
3.9371 | -84.0916
3.9374 | -84.094
3.9377 | -84.0951
3.9381 | -84.0961
3.9384 | -84.0952
3.9388 | -84.0953
3.9391 | -84.0961
3.9394 | -84.0963
3.9398 | -84.0963
3.9401 | -84.0959
3.9405 | -84.0993
3.9408 | -84.1023
3.9411 | -84.1028
3.9415 | -84.1028
3.9418 | -84.1037
3.9422 | -84.105
3.9425 | -84.1056
3.9429 | -84.1062
3.9432 | -84.1073
3.9435 | -84.1081
3.9439 | -84.1087
3.9442 | -84.1091
3.9446 | -84.1092
3.9449 | -84.1097
3.9452 | -84.1099
3.9456 | -84.1104
3.9459 | -84.1102
3.9463 | -84.1107
3.9466 | -84.1117
3.9469 | -84.1136
3.9473 | -84.1147
3.9476 | -84.1154
3.948 | -84.1167
3.9483 | -84.1184
3.9486 | -84.1209
3.949 | -84.1215
3.9493 | -84.1224
3.9497 | -84.1235
3.95 | -84.1241
3.9503 | -84.1215
3.9507 | -84.1191
3.951 | -84.1199
3.9514 | -84.1206
3.9517 | -84.1211
3.952 | -84.1226
3.9524 | -84.1248
3.9527 | -84.1262
3.9531 | -84.1246
3.9534 | -84.1221
3.9537 | -84.1198
3.9541 | -84.1187
3.9544 | -84.1196
3.9548 | -84.1188
3.9551 | -84.1185
3.9555 | -84.1186
3.9558 | -84.1187
3.9561 | -84.119
3.9565 | -84.1195
3.9568 | -84.1183
3.9572 | -84.1143
3.9575 | -84.1131
3.9578 | -84.1142
3.9582 | -84.1143
3.9585 | -84.1151
3.9589 | -84.1141
3.9592 | -84.1139
3.9595 | -84.1121
3.9599 | -84.112
3.9602 | -84.1124
3.9606 | -84.1132
3.9609 | -84.1142
3.9612 | -84.1151
3.9616 | -84.116
3.9619 | -84.1172
3.9623 | -84.118
3.9626 | -84.1201
3.9629 | -84.1223
3.9633 | -84.1225
3.9636 | -84.1228
3.964 | -84.124
3.9643 | -84.1254
3.9646 | -84.1261
3.965 | -84.1267
3.9653 | -84.127
3.9657 | -84.1275
3.966 | -84.1281
3.9663 | -84.1289
3.9667 | -84.1294
3.967 | -84.1298
3.9674 | -84.1304
3.9677 | -84.1317
3.968 | -84.1339
3.9684 | -84.136
3.9687 | -84.137
3.9691 | -84.1377
3.9694 | -84.137
3.9698 | -84.1377
3.9701 | -84.1383
3.9704 | -84.1381
3.9708 | -84.1393
3.9711 | -84.1403
3.9715 | -84.1402
3.9718 | -84.14
3.9721 | -84.1411
3.9725 | -84.1418
3.9728 | -84.1414
3.9732 | -84.1431
3.9735 | -84.1457
3.9738 | -84.1461
3.9742 | -84.1446
3.9745 | -84.1453
3.9749 | -84.1447
3.9752 | -84.1448
3.9755 | -84.1452
3.9759 | -84.1455
3.9762 | -84.146
3.9766 | -84.1473
3.9769 | -84.148
3.9772 | -84.1484
3.9776 | -84.1493
3.9779 | -84.1493
3.9783 | -84.1478
3.9786 | -84.1469
3.9789 | -84.1459
3.9793 | -84.146
3.9796 | -84.1462
3.98 | -84.1468
3.9803 | -84.1477
3.9806 | -84.148
3.981 | -84.1474
3.9813 | -84.1461
3.9817 | -84.1468
3.982 | -84.1466
3.9823 | -84.1462
3.9827 | -84.1467
3.983 | -84.1466
3.9834 | -84.1467
3.9837 | -84.147
3.9841 | -84.1476
3.9844 | -84.1482
3.9847 | -84.1483
3.9851 | -84.1497
3.9854 | -84.1507
3.9858 | -84.1517
3.9861 | -84.1527
3.9864 | -84.1536
3.9868 | -84.1547
3.9871 | -84.1552
3.9875 | -84.1563
3.9878 | -84.1573
3.9881 | -84.1579
3.9885 | -84.1583
3.9888 | -84.1583
3.9892 | -84.1578
3.9895 | -84.1579
3.9898 | -84.1587
3.9902 | -84.1586
3.9905 | -84.1581
3.9909 | -84.1575
3.9912 | -84.1573
3.9915 | -84.1574
3.9919 | -84.1563
3.9922 | -84.1552
3.9926 | -84.1551
3.9929 | -84.1526
3.9932 | -84.1507
3.9936 | -84.151
3.9939 | -84.1508
3.9943 | -84.149
3.9946 | -84.146
3.9949 | -84.1445
3.9953 | -84.1441
3.9956 | -84.1436
3.996 | -84.1429
3.9963 | -84.1429
3.9966 | -84.1423
3.997 | -84.1416
3.9973 | -84.1413
3.9977 | -84.1415
3.998 | -84.1395
3.9984 | -84.1404
3.9987 | -84.1418
3.999 | -84.1428
3.9994 | -84.1427
3.9997 | -84.1425
4.0001 | -84.1427
4.0004 | -84.1421
4.0007 | -84.1424
4.0011 | -84.1431
4.0014 | -84.1438
4.0018 | -84.1441
4.0021 | -84.1437
4.0024 | -84.1437
4.0028 | -84.1434
4.0031 | -84.1432
4.0035 | -84.1435
4.0038 | -84.1433
4.0041 | -84.1435
4.0045 | -84.1429
4.0048 | -84.1427
4.0052 | -84.1416
4.0055 | -84.1404
4.0058 | -84.1399
4.0062 | -84.1385
4.0065 | -84.1386
4.0069 | -84.1399
4.0072 | -84.1398
4.0075 | -84.139
4.0079 | -84.138
4.0082 | -84.1381
4.0086 | -84.1374
4.0089 | -84.1373
4.0092 | -84.1365
4.0096 | -84.1358
4.0099 | -84.1351
4.0103 | -84.1346
4.0106 | -84.1351
4.011 | -84.1351
4.0113 | -84.1345
4.0116 | -84.1342
4.012 | -84.1337
4.0123 | -84.1334
4.0127 | -84.1339
4.013 | -84.1336
4.0133 | -84.1328
4.0137 | -84.1322
4.014 | -84.1317
4.0144 | -84.1306
4.0147 | -84.1302
4.015 | -84.1298
4.0154 | -84.1291
4.0157 | -84.1279
4.0161 | -84.1272
4.0164 | -84.1283
4.0167 | -84.1295
4.0171 | -84.1305
4.0174 | -84.1312
4.0178 | -84.131
4.0181 | -84.1285
4.0184 | -84.1277
4.0188 | -84.126
4.0191 | -84.1246
4.0195 | -84.1248
4.0198 | -84.1235
4.0201 | -84.1198
4.0205 | -84.1187
4.0208 | -84.1184
4.0212 | -84.1183
4.0215 | -84.1178
4.0218 | -84.1186
4.0222 | -84.119
4.0225 | -84.1165
4.0229 | -84.1158
4.0232 | -84.1166
4.0235 | -84.1175
4.0239 | -84.117
4.0242 | -84.1158
4.0246 | -84.1156
4.0249 | -84.1146
4.0253 | -84.1142
4.0256 | -84.1137
4.0259 | -84.1131
4.0263 | -84.1126
4.0266 | -84.1116
4.027 | -84.11
4.0273 | -84.109
4.0276 | -84.1087
4.028 | -84.109
4.0283 | -84.1094
4.0287 | -84.1082
4.029 | -84.1071
4.0293 | -84.1071
4.0297 | -84.1079
4.03 | -84.1073
4.0304 | -84.1061
4.0307 | -84.104
4.031 | -84.1019
4.0314 | -84.0999
4.0317 | -84.0986
4.0321 | -84.0975
4.0324 | -84.0975
4.0327 | -84.0974
4.0331 | -84.0963
4.0334 | -84.0952
4.0338 | -84.0942
4.0341 | -84.0934
4.0344 | -84.0931
4.0348 | -84.0927
4.0351 | -84.0917
4.0355 | -84.091
4.0358 | -84.0907
4.0361 | -84.0908
4.0365 | -84.0903
4.0368 | -84.0891
4.0372 | -84.088
4.0375 | -84.0874
4.0378 | -84.0869
4.0382 | -84.086
4.0385 | -84.0851
4.0389 | -84.0836
4.0392 | -84.0815
4.0396 | -84.0809
4.0399 | -84.0809
4.0402 | -84.0797
4.0406 | -84.0778
4.0409 | -84.0766
4.0413 | -84.0751
4.0416 | -84.0731
4.0419 | -84.072
4.0423 | -84.0712
4.0426 | -84.0696
4.043 | -84.0709
4.0433 | -84.0721
4.0436 | -84.0738
4.044 | -84.0729
4.0443 | -84.0718
4.0447 | -84.0709
4.045 | -84.0702
4.0453 | -84.0687
4.0457 | -84.0672
4.046 | -84.0678
4.0464 | -84.0678
4.0467 | -84.0673
4.047 | -84.0654
4.0474 | -84.0633
4.0477 | -84.0619
4.0481 | -84.0603
4.0484 | -84.059
4.0487 | -84.0579
4.0491 | -84.0565
4.0494 | -84.055
4.0498 | -84.054
4.0501 | -84.0521
4.0504 | -84.0486
4.0508 | -84.0486
4.0511 | -84.0484
4.0515 | -84.047
4.0518 | -84.0463
4.0521 | -84.0454
4.0525 | -84.0444
4.0528 | -84.0434
4.0532 | -84.0423
4.0535 | -84.0416
4.0539 | -84.0404
4.0542 | -84.0382
4.0545 | -84.0369
4.0549 | -84.0369
4.0552 | -84.0364
4.0556 | -84.0352
4.0559 | -84.0347
4.0562 | -84.0345
4.0566 | -84.0343
4.0569 | -84.0344
4.0573 | -84.0339
4.0576 | -84.0333
4.0579 | -84.0335
4.0583 | -84.034
4.0586 | -84.0335
4.059 | -84.0312
4.0593 | -84.0274
4.0596 | -84.0244
4.06 | -84.0221
4.0603 | -84.0203
4.0607 | -84.0191
4.061 | -84.0186
4.0613 | -84.0172
4.0617 | -84.0157
4.062 | -84.0153
4.0624 | -84.0143
4.0627 | -84.0131
4.063 | -84.0118
4.0634 | -84.01
4.0637 | -84.0095
4.0641 | -84.0088
4.0644 | -84.0067
4.0647 | -84.0046
4.0651 | -84.0029
4.0654 | -84.0018
4.0658 | -83.9997
4.0661 | -83.9979
4.0664 | -83.9973
4.0668 | -83.9971
4.0671 | -83.9965
4.0675 | -83.9952
4.0678 | -83.9948
4.0682 | -83.9952
4.0685 | -83.9948
4.0688 | -83.9938
4.0692 | -83.9933
4.0695 | -83.9937
4.0699 | -83.991
4.0702 | -83.9895
4.0705 | -83.9895
4.0709 | -83.9895
4.0712 | -83.9892
4.0716 | -83.9885
4.0719 | -83.9876
4.0722 | -83.987
4.0726 | -83.9865
4.0729 | -83.9862
4.0733 | -83.9854
4.0736 | -83.984
4.0739 | -83.983
4.0743 | -83.982
4.0746 | -83.9816
4.075 | -83.9819
4.0753 | -83.9827
4.0756 | -83.9834
4.076 | -83.9837
4.0763 | -83.9835
4.0767 | -83.9842
4.077 | -83.9838
4.0773 | -83.983
4.0777 | -83.9813
4.078 | -83.9792
4.0784 | -83.9768
4.0787 | -83.9749
4.079 | -83.9733
4.0794 | -83.9729
4.0797 | -83.9723
4.0801 | -83.9719
4.0804 | -83.9703
4.0808 | -83.9694
4.0811 | -83.9697
4.0814 | -83.9706
4.0818 | -83.9709
4.0821 | -83.9698
4.0825 | -83.969
4.0828 | -83.9684
4.0831 | -83.9688
4.0835 | -83.9682
4.0838 | -83.9674
4.0842 | -83.9667
4.0845 | -83.9659
4.0848 | -83.9662
4.0852 | -83.9663
4.0855 | -83.9658
4.0859 | -83.966
4.0862 | -83.9665
4.0865 | -83.967
4.0869 | -83.9664
4.0872 | -83.9658
4.0876 | -83.9664
4.0879 | -83.9663
4.0882 | -83.9657
4.0886 | -83.9655
4.0889 | -83.9652
4.0893 | -83.9644
4.0896 | -83.9641
4.0899 | -83.9633
4.0903 | -83.9627
4.0906 | -83.9627
4.091 | -83.9622
4.0913 | -83.9621
4.0916 | -83.9615
4.092 | -83.9621
4.0923 | -83.9639
4.0927 | -83.9645
4.093 | -83.9648
4.0933 | -83.964
4.0937 | -83.9629
4.094 | -83.9615
4.0944 | -83.9609
4.0947 | -83.9603
4.0951 | -83.9597
4.0954 | -83.9594
4.0957 | -83.9592
4.0961 | -83.9591
4.0964 | -83.9605
4.0968 | -83.9598
4.0971 | -83.958
4.0974 | -83.9573
4.0978 | -83.9572
4.0981 | -83.9572
4.0985 | -83.9569
4.0988 | -83.9563
4.0991 | -83.9567
4.0995 | -83.956
4.0998 | -83.9555
4.1002 | -83.9555
4.1005 | -83.9554
4.1008 | -83.9552
4.1012 | -83.9551
4.1015 | -83.9541
4.1019 | -83.9534
4.1022 | -83.9531
4.1025 | -83.9532
4.1029 | -83.952
4.1032 | -83.9511
4.1036 | -83.9508
4.1039 | -83.9516
4.1042 | -83.9521
4.1046 | -83.952
4.1049 | -83.9514
4.1053 | -83.9511
4.1056 | -83.9506
4.1059 | -83.9511
4.1063 | -83.9517
4.1066 | -83.952
4.107 | -83.9525
4.1073 | -83.9532
4.1076 | -83.9527
4.108 | -83.9506
4.1083 | -83.9485
4.1087 | -83.9485
4.109 | -83.9489
4.1094 | -83.9504
4.1097 | -83.9505
4.11 | -83.9503
4.1104 | -83.9505
4.1107 | -83.9504
4.1111 | -83.9504
4.1114 | -83.9504
4.1117 | -83.9506
4.1121 | -83.9509
4.1124 | -83.9513
4.1128 | -83.9513
4.1131 | -83.952
4.1134 | -83.9543
4.1138 | -83.956
4.1141 | -83.9571
4.1145 | -83.9588
4.1148 | -83.959
4.1151 | -83.958
4.1155 | -83.9564
4.1158 | -83.9557
4.1162 | -83.9555
4.1165 | -83.9548
4.1168 | -83.9549
4.1172 | -83.9555
4.1175 | -83.9559
4.1179 | -83.9557
4.1182 | -83.956
4.1185 | -83.9558
4.1189 | -83.9549
4.1192 | -83.955
4.1196 | -83.9548
4.1199 | -83.9547
4.1202 | -83.9552
4.1206 | -83.9561
4.1209 | -83.9567
4.1213 | -83.9571
4.1216 | -83.9576
4.1219 | -83.9566
4.1223 | -83.954
4.1226 | -83.9494
4.123 | -83.946
4.1233 | -83.9471
4.1237 | -83.9497
4.124 | -83.9524
4.1243 | -83.9522
4.1247 | -83.9541
4.125 | -83.9558
4.1254 | -83.9556
4.1257 | -83.955
4.126 | -83.9548
4.1264 | -83.9548
4.1267 | -83.9558
4.1271 | -83.9571
4.1274 | -83.9574
4.1277 | -83.9569
4.1281 | -83.9569
4.1284 | -83.9576
4.1288 | -83.9583
4.1291 | -83.9594
4.1294 | -83.9601
4.1298 | -83.9607
4.1301 | -83.9609
4.1305 | -83.9606
4.1308 | -83.9604
4.1311 | -83.9599
4.1315 | -83.9606
4.1318 | -83.9618
4.1322 | -83.9621
4.1325 | -83.9619
4.1328 | -83.9615
4.1332 | -83.9614
4.1335 | -83.961
4.1339 | -83.9607
4.1342 | -83.9619
4.1345 | -83.9626
4.1349 | -83.9632
4.1352 | -83.9636
4.1356 | -83.9642
4.1359 | -83.9642
4.1363 | -83.9647
4.1366 | -83.9655
4.1369 | -83.9655
4.1373 | -83.9658
4.1376 | -83.9663
4.138 | -83.9677
4.1383 | -83.9687
4.1386 | -83.9689
4.139 | -83.9684
4.1393 | -83.968
4.1397 | -83.9683
4.14 | -83.9688
4.1403 | -83.9697
4.1407 | -83.9709
4.141 | -83.9712
4.1414 | -83.9701
4.1417 | -83.9691
4.142 | -83.9682
4.1424 | -83.9689
4.1427 | -83.9688
4.1431 | -83.9679
4.1434 | -83.9656
4.1437 | -83.9639
4.1441 | -83.9626
4.1444 | -83.9624
4.1448 | -83.9635
4.1451 | -83.9634
4.1454 | -83.9645
4.1458 | -83.9644
4.1461 | -83.9633
4.1465 | -83.9628
4.1468 | -83.9636
4.1471 | -83.9639
4.1475 | -83.9627
4.1478 | -83.962
4.1482 | -83.962
4.1485 | -83.9625
4.1488 | -83.963
4.1492 | -83.9637
4.1495 | -83.9634
4.1499 | -83.9633
4.1502 | -83.9622
4.1505 | -83.9608
4.1509 | -83.9608
4.1512 | -83.9619
4.1516 | -83.9624
4.1519 | -83.9623
4.1523 | -83.9622
4.1526 | -83.9631
4.1529 | -83.963
4.1533 | -83.9632
4.1536 | -83.9632
4.154 | -83.9626
4.1543 | -83.9617
4.1546 | -83.961
4.155 | -83.9589
4.1553 | -83.9583
4.1557 | -83.9583
4.156 | -83.9577
4.1563 | -83.9576
4.1567 | -83.9575
4.157 | -83.9574
4.1574 | -83.9557
4.1577 | -83.955
4.158 | -83.9551
4.1584 | -83.9556
4.1587 | -83.9559
4.1591 | -83.9562
4.1594 | -83.956
4.1597 | -83.9555
4.1601 | -83.9563
4.1604 | -83.9575
4.1608 | -83.957
4.1611 | -83.9567
4.1614 | -83.9575
4.1618 | -83.9584
4.1621 | -83.9589
4.1625 | -83.9593
4.1628 | -83.9601
4.1631 | -83.9604
4.1635 | -83.9596
4.1638 | -83.9597
4.1642 | -83.9592
4.1645 | -83.9594
4.1649 | -83.9602
4.1652 | -83.9594
4.1655 | -83.9583
4.1659 | -83.9581
4.1662 | -83.9582
4.1666 | -83.9577
4.1669 | -83.9568
4.1672 | -83.9577
4.1676 | -83.9579
4.1679 | -83.959
4.1683 | -83.9614
4.1686 | -83.9621
4.1689 | -83.962
4.1693 | -83.9615
4.1696 | -83.9613
4.17 | -83.9611
4.1703 | -83.9615
4.1706 | -83.962
4.171 | -83.962
4.1713 | -83.9619
4.1717 | -83.9613
4.172 | -83.9604
4.1723 | -83.9592
4.1727 | -83.9581
4.173 | -83.9582
4.1734 | -83.9581
4.1737 | -83.958
4.174 | -83.9579
4.1744 | -83.9578
4.1747 | -83.9579
4.1751 | -83.9567
4.1754 | -83.9549
4.1757 | -83.9543
4.1761 | -83.9539
4.1764 | -83.9535
4.1768 | -83.9533
4.1771 | -83.9543
4.1774 | -83.9547
4.1778 | -83.9545
4.1781 | -83.9553
4.1785 | -83.9554
4.1788 | -83.9552
4.1792 | -83.9552
4.1795 | -83.9554
4.1798 | -83.9566
4.1802 | -83.9575
4.1805 | -83.9571
4.1809 | -83.9561
4.1812 | -83.9544
4.1815 | -83.9534
4.1819 | -83.9531
4.1822 | -83.953
4.1826 | -83.9525
4.1829 | -83.9522
4.1832 | -83.952
4.1836 | -83.9517
4.1839 | -83.952
4.1843 | -83.9522
4.1846 | -83.9519
4.1849 | -83.9516
4.1853 | -83.9511
4.1856 | -83.9505
4.186 | -83.9497
4.1863 | -83.9493
4.1866 | -83.9492
4.187 | -83.9493
4.1873 | -83.9491
4.1877 | -83.9495
4.188 | -83.9497
4.1883 | -83.9492
4.1887 | -83.949
4.189 | -83.9489
4.1894 | -83.9487
4.1897 | -83.9488
4.19 | -83.948
4.1904 | -83.9473
4.1907 | -83.9471
4.1911 | -83.9467
4.1914 | -83.9465
4.1917 | -83.9469
4.1921 | -83.9464
4.1924 | -83.9464
4.1928 | -83.9464
4.1931 | -83.9463
4.1935 | -83.9458
4.1938 | -83.9457
4.1941 | -83.9457
4.1945 | -83.9457
4.1948 | -83.9446
4.1952 | -83.9426
4.1955 | -83.9413
4.1958 | -83.9409
4.1962 | -83.9411
4.1965 | -83.9416
4.1969 | -83.941
4.1972 | -83.9405
4.1975 | -83.9392
4.1979 | -83.9372
4.1982 | -83.9362
4.1986 | -83.9357
4.1989 | -83.9344
4.1992 | -83.9337
4.1996 | -83.933
4.1999 | -83.9315
4.2003 | -83.9321
4.2006 | -83.934
4.2009 | -83.9334
4.2013 | -83.9325
4.2016 | -83.932
4.202 | -83.9323
4.2023 | -83.9324
4.2026 | -83.9317
4.203 | -83.9299
4.2033 | -83.9276
4.2037 | -83.926
4.204 | -83.9258
4.2043 | -83.9234
4.2047 | -83.9234
4.205 | -83.9231
4.2054 | -83.9209
4.2057 | -83.9195
4.2061 | -83.9192
4.2064 | -83.9196
4.2067 | -83.9216
4.2071 | -83.9244
4.2074 | -83.9246
4.2078 | -83.9234
4.2081 | -83.9224
4.2084 | -83.9214
4.2088 | -83.9202
4.2091 | -83.9205
4.2095 | -83.9207
4.2098 | -83.9199
4.2101 | -83.9192
4.2105 | -83.9184
4.2108 | -83.9182
4.2112 | -83.9179
4.2115 | -83.9171
4.2118 | -83.917
4.2122 | -83.9171
4.2125 | -83.9167
4.2129 | -83.916
4.2132 | -83.9154
4.2135 | -83.9147
4.2139 | -83.9146
4.2142 | -83.9143
4.2146 | -83.9142
4.2149 | -83.9129
4.2152 | -83.9115
4.2156 | -83.9105
4.2159 | -83.9095
4.2163 | -83.9088
4.2166 | -83.9084
4.2169 | -83.9084
4.2173 | -83.9077
4.2176 | -83.907
4.218 | -83.9059
4.2183 | -83.9047
4.2186 | -83.9041
4.219 | -83.9035
4.2193 | -83.9027
4.2197 | -83.9016
4.22 | -83.9017
4.2203 | -83.9013
4.2207 | -83.9007
4.221 | -83.9002
4.2214 | -83.8993
4.2217 | -83.8988
4.2221 | -83.8972
4.2224 | -83.8958
4.2227 | -83.8949
4.2231 | -83.8945
4.2234 | -83.8939
4.2238 | -83.893
4.2241 | -83.8894
4.2244 | -83.8883
4.2248 | -83.8874
4.2251 | -83.8866
4.2255 | -83.886
4.2258 | -83.8846
4.2261 | -83.8839
4.2265 | -83.8827
4.2268 | -83.8812
4.2272 | -83.8795
4.2275 | -83.8782
4.2278 | -83.8774
4.2282 | -83.877
4.2285 | -83.876
4.2289 | -83.8751
4.2292 | -83.8738
4.2295 | -83.8731
4.2299 | -83.8729
4.2302 | -83.8718
4.2306 | -83.8708
4.2309 | -83.87
4.2312 | -83.8687
4.2316 | -83.8671
4.2319 | -83.8658
4.2323 | -83.8644
4.2326 | -83.8633
4.2329 | -83.8634
4.2333 | -83.8624
4.2336 | -83.8608
4.234 | -83.8594
4.2343 | -83.8583
4.2347 | -83.8579
4.235 | -83.8598
4.2353 | -83.8613
4.2357 | -83.86
4.236 | -83.8593
4.2364 | -83.8583
4.2367 | -83.8572
4.237 | -83.8563
4.2374 | -83.8557
4.2377 | -83.8577
4.2381 | -83.8596
4.2384 | -83.8619
4.2387 | -83.8622
4.2391 | -83.8611
4.2394 | -83.8619
4.2398 | -83.862
4.2401 | -83.8613
4.2404 | -83.8606
4.2408 | -83.8598
4.2411 | -83.8591
4.2415 | -83.8601
4.2418 | -83.863
4.2421 | -83.8635
4.2425 | -83.8628
4.2428 | -83.8617
4.2432 | -83.8602
4.2435 | -83.8601
4.2438 | -83.8596
4.2442 | -83.8584
4.2445 | -83.858
4.2449 | -83.8591
4.2452 | -83.8586
4.2455 | -83.8576
4.2459 | -83.8568
4.2462 | -83.8563
4.2466 | -83.8553
4.2469 | -83.8547
4.2472 | -83.8542
4.2476 | -83.8535
4.2479 | -83.8521
4.2483 | -83.8508
4.2486 | -83.8507
4.249 | -83.8499
4.2493 | -83.8491
4.2496 | -83.8486
4.25 | -83.8482
4.2503 | -83.8478
4.2507 | -83.8475
4.251 | -83.8466
4.2513 | -83.8456
4.2517 | -83.8452
4.252 | -83.8444
4.2524 | -83.8436
4.2527 | -83.8413
4.253 | -83.8399
4.2534 | -83.8394
4.2537 | -83.8388
4.2541 | -83.8385
4.2544 | -83.8379
4.2547 | -83.8372
4.2551 | -83.8366
4.2554 | -83.8373
4.2558 | -83.8359
4.2561 | -83.8331
4.2564 | -83.8326
4.2568 | -83.8315
4.2571 | -83.8302
4.2575 | -83.829
4.2578 | -83.8285
4.2581 | -83.8281
4.2585 | -83.8278
4.2588 | -83.8285
4.2592 | -83.8291
4.2595 | -83.828
4.2598 | -83.8276
4.2602 | -83.828
4.2605 | -83.8268
4.2609 | -83.8269
4.2612 | -83.8264
4.2615 | -83.8256
4.2619 | -83.8254
4.2622 | -83.8253
4.2626 | -83.8253
4.2629 | -83.8254
4.2633 | -83.8258
4.2636 | -83.826
4.2639 | -83.8259
4.2643 | -83.8258
4.2646 | -83.8265
4.265 | -83.8263
4.2653 | -83.8276
4.2656 | -83.8282
4.266 | -83.8282
4.2663 | -83.8282
4.2667 | -83.8281
4.267 | -83.8279
4.2673 | -83.8273
4.2677 | -83.8275
4.268 | -83.8283
4.2684 | -83.8282
4.2687 | -83.8288
4.269 | -83.8294
4.2694 | -83.8291
4.2697 | -83.8292
4.2701 | -83.8292
4.2704 | -83.8286
4.2707 | -83.8283
4.2711 | -83.8287
4.2714 | -83.8296
4.2718 | -83.8301
4.2721 | -83.8293
4.2724 | -83.8275
4.2728 | -83.8257
4.2731 | -83.827
4.2735 | -83.8262
4.2738 | -83.8255
4.2741 | -83.8258
4.2745 | -83.8261
4.2748 | -83.8263
4.2752 | -83.8265
4.2755 | -83.8264
4.2759 | -83.8262
4.2762 | -83.8262
4.2765 | -83.8267
4.2769 | -83.826
4.2772 | -83.8244
4.2776 | -83.824
4.2779 | -83.8245
4.2782 | -83.8255
4.2786 | -83.826
4.2789 | -83.8265
4.2793 | -83.828
4.2796 | -83.8297
4.2799 | -83.8301
4.2803 | -83.8319
4.2806 | -83.8333
4.281 | -83.8337
4.2813 | -83.8341
4.2816 | -83.8362
4.282 | -83.8389
4.2823 | -83.8402
4.2827 | -83.8404
4.283 | -83.8401
4.2833 | -83.8404
4.2837 | -83.841
4.284 | -83.841
4.2844 | -83.8412
4.2847 | -83.8418
4.285 | -83.8417
4.2854 | -83.8421
4.2857 | -83.8426
4.2861 | -83.8441
4.2864 | -83.8452
4.2867 | -83.8456
4.2871 | -83.8458
4.2874 | -83.8466
4.2878 | -83.8471
4.2881 | -83.8474
4.2884 | -83.8479
4.2888 | -83.848
4.2891 | -83.8477
4.2895 | -83.8469
4.2898 | -83.8474
4.2901 | -83.8471
4.2905 | -83.8461
4.2908 | -83.8466
4.2912 | -83.8472
4.2915 | -83.848
4.2919 | -83.8486
4.2922 | -83.8488
4.2925 | -83.8497
4.2929 | -83.8522
4.2932 | -83.8525
4.2936 | -83.8531
4.2939 | -83.8539
4.2942 | -83.8545
4.2946 | -83.8558
4.2949 | -83.8564
4.2953 | -83.8572
4.2956 | -83.8585
4.2959 | -83.8599
4.2963 | -83.8612
4.2966 | -83.8604
4.297 | -83.8603
4.2973 | -83.8612
4.2976 | -83.8615
4.298 | -83.8624
4.2983 | -83.8634
4.2987 | -83.8646
4.299 | -83.8659
4.2993 | -83.8667
4.2997 | -83.8665
4.3 | -83.8667
4.3004 | -83.8672
4.3007 | -83.8671
4.301 | -83.8672
4.3014 | -83.8678
4.3017 | -83.8687
4.3021 | -83.8702
4.3024 | -83.8715
4.3027 | -83.874
4.3031 | -83.875
4.3034 | -83.8763
4.3038 | -83.8773
4.3041 | -83.878
4.3045 | -83.879
4.3048 | -83.8808
4.3051 | -83.8828
4.3055 | -83.8844
4.3058 | -83.8855
4.3062 | -83.8867
4.3065 | -83.8877
4.3068 | -83.8887
4.3072 | -83.8901
4.3075 | -83.893
4.3079 | -83.8946
4.3082 | -83.8968
4.3085 | -83.899
4.3089 | -83.8989
4.3092 | -83.8997
4.3096 | -83.903
4.3099 | -83.904
4.3102 | -83.905
4.3106 | -83.9049
4.3109 | -83.9048
4.3113 | -83.9057
4.3116 | -83.9074
4.3119 | -83.9097
4.3123 | -83.9117
4.3126 | -83.9127
4.313 | -83.9136
4.3133 | -83.9144
4.3136 | -83.9158
4.314 | -83.9174
4.3143 | -83.9181
4.3147 | -83.9183
4.315 | -83.9187
4.3153 | -83.9192
4.3157 | -83.9192
4.316 | -83.9196
4.3164 | -83.9213
4.3167 | -83.9227
4.317 | -83.9233
4.3174 | -83.924
4.3177 | -83.9251
4.3181 | -83.9262
4.3184 | -83.9268
4.3188 | -83.9288
4.3191 | -83.9304
4.3194 | -83.9317
4.3198 | -83.9322
4.3201 | -83.9327
4.3205 | -83.9337
4.3208 | -83.9354
4.3211 | -83.9364
4.3215 | -83.9374
4.3218 | -83.9386
4.3222 | -83.9398
4.3225 | -83.9403
4.3228 | -83.9415
4.3232 | -83.942
4.3235 | -83.9433
4.3239 | -83.9441
4.3242 | -83.9442
4.3245 | -83.9446
4.3249 | -83.9433
4.3252 | -83.9412
4.3256 | -83.9406
4.3259 | -83.9415
4.3262 | -83.9426
4.3266 | -83.9437
4.3269 | -83.9445
4.3273 | -83.946
4.3276 | -83.9476
4.3279 | -83.9489
4.3283 | -83.9484
4.3286 | -83.9483
4.329 | -83.949
4.3293 | -83.9495
4.3296 | -83.9508
4.33 | -83.9529
4.3303 | -83.954
4.3307 | -83.9548
4.331 | -83.9559
4.3313 | -83.9574
4.3317 | -83.9595
4.332 | -83.9608
4.3324 | -83.9622
4.3327 | -83.9632
4.3331 | -83.9631
4.3334 | -83.9635
4.3337 | -83.966
4.3341 | -83.9666
4.3344 | -83.9677
4.3348 | -83.9688
4.3351 | -83.9698
4.3354 | -83.9704
4.3358 | -83.9716
4.3361 | -83.9718
4.3365 | -83.9728
4.3368 | -83.9736
4.3371 | -83.9737
4.3375 | -83.9737
4.3378 | -83.9749
4.3382 | -83.9772
4.3385 | -83.98
4.3388 | -83.9809
4.3392 | -83.981
4.3395 | -83.9813
4.3399 | -83.981
4.3402 | -83.9803
4.3405 | -83.9802
4.3409 | -83.9803
4.3412 | -83.9802
4.3416 | -83.9815
4.3419 | -83.9818
4.3422 | -83.9845
4.3426 | -83.9856
4.3429 | -83.986
4.3433 | -83.9857
4.3436 | -83.9846
4.3439 | -83.9831
4.3443 | -83.983
4.3446 | -83.9842
4.345 | -83.9845
4.3453 | -83.9849
4.3457 | -83.9863
4.346 | -83.988
4.3463 | -83.988
4.3467 | -83.9879
4.347 | -83.9883
4.3474 | -83.9889
4.3477 | -83.9885
4.348 | -83.9887
4.3484 | -83.9893
4.3487 | -83.9895
4.3491 | -83.9893
4.3494 | -83.9899
4.3497 | -83.9901
4.3501 | -83.9895
4.3504 | -83.99
4.3508 | -83.9907
4.3511 | -83.9927
4.3514 | -83.9959
4.3518 | -83.9963
4.3521 | -83.9991
4.3525 | -84.0002
4.3528 | -83.9997
4.3531 | -83.9998
4.3535 | -84.0001
4.3538 | -84.0009
4.3542 | -84.0019
4.3545 | -84.0023
4.3548 | -84.0026
4.3552 | -84.0021
4.3555 | -84.0016
4.3559 | -84.0019
4.3562 | -84.0008
4.3565 | -83.9989
4.3569 | -83.9989
4.3572 | -84.0003
4.3576 | -84.0014
4.3579 | -84.0018
4.3582 | -84.003
4.3586 | -84.0035
4.3589 | -84.0035
4.3593 | -84.0036
4.3596 | -84.0039
4.36 | -84.0042
4.3603 | -84.0045
4.3606 | -84.0048
4.361 | -84.005
4.3613 | -84.0053
4.3617 | -84.0048
4.362 | -84.005
4.3623 | -84.0054
4.3627 | -84.0058
4.363 | -84.0079
4.3634 | -84.009
4.3637 | -84.009
4.364 | -84.009
4.3644 | -84.009
4.3647 | -84.0089
4.3651 | -84.0081
4.3654 | -84.0072
4.3657 | -84.0072
4.3661 | -84.0073
4.3664 | -84.0071
4.3668 | -84.0085
4.3671 | -84.0094
4.3674 | -84.0093
4.3678 | -84.0091
4.3681 | -84.0087
4.3685 | -84.0085
4.3688 | -84.0087
4.3691 | -84.0078
4.3695 | -84.0071
4.3698 | -84.0064
4.3702 | -84.0046
4.3705 | -84.0029
4.3708 | -84.0016
4.3712 | -84.0012
4.3715 | -83.9995
4.3719 | -84.0005
4.3722 | -84.0018
4.3725 | -84.0028
4.3729 | -84.0034
4.3732 | -84.0032
4.3736 | -84.0022
4.3739 | -84.0017
4.3743 | -84.0025
4.3746 | -84.002
4.3749 | -84.0014
4.3753 | -84.0011
4.3756 | -84.0001
4.376 | -83.9994
4.3763 | -83.9991
4.3766 | -83.9985
4.377 | -83.998
4.3773 | -83.9976
4.3777 | -83.9972
4.378 | -83.9969
4.3783 | -83.9967
4.3787 | -83.9965
4.379 | -83.9962
4.3794 | -83.9959
4.3797 | -83.9953
4.38 | -83.995
4.3804 | -83.9938
4.3807 | -83.9928
4.3811 | -83.9922
4.3814 | -83.9917
4.3817 | -83.9915
4.3821 | -83.9909
4.3824 | -83.9904
4.3828 | -83.9896
4.3831 | -83.9891
4.3834 | -83.9885
4.3838 | -83.9886
4.3841 | -83.9879
4.3845 | -83.9869
4.3848 | -83.9862
4.3851 | -83.9854
4.3855 | -83.9851
4.3858 | -83.985
4.3862 | -83.9846
4.3865 | -83.9835
4.3868 | -83.9826
4.3872 | -83.9816
4.3875 | -83.9811
4.3879 | -83.9804
4.3882 | -83.9799
4.3886 | -83.9805
4.3889 | -83.9804
4.3892 | -83.9802
4.3896 | -83.9799
4.3899 | -83.9787
4.3903 | -83.9775
4.3906 | -83.9767
4.3909 | -83.9747
4.3913 | -83.9737
4.3916 | -83.973
4.392 | -83.9727
4.3923 | -83.9735
4.3926 | -83.9736
4.393 | -83.9724
4.3933 | -83.9705
4.3937 | -83.9686
4.394 | -83.9687
4.3943 | -83.97
4.3947 | -83.9709
4.395 | -83.9701
4.3954 | -83.9686
4.3957 | -83.9672
4.396 | -83.9656
4.3964 | -83.9647
4.3967 | -83.9649
4.3971 | -83.9649
4.3974 | -83.964
4.3977 | -83.9633
4.3981 | -83.9635
4.3984 | -83.9633
4.3988 | -83.9629
4.3991 | -83.9616
4.3994 | -83.9606
4.3998 | -83.96
4.4001 | -83.9596
4.4005 | -83.9583
4.4008 | -83.9559
4.4011 | -83.9551
4.4015 | -83.9541
4.4018 | -83.9533
4.4022 | -83.9531
4.4025 | -83.9526
4.4029 | -83.9519
4.4032 | -83.9509
4.4035 | -83.9504
4.4039 | -83.9508
4.4042 | -83.9513
4.4046 | -83.9508
4.4049 | -83.9503
4.4052 | -83.9489
4.4056 | -83.9473
4.4059 | -83.9439
4.4063 | -83.9409
4.4066 | -83.9396
4.4069 | -83.9385
4.4073 | -83.9371
4.4076 | -83.9358
4.408 | -83.9341
4.4083 | -83.9325
4.4086 | -83.9318
4.409 | -83.9308
4.4093 | -83.93
4.4097 | -83.9281
4.41 | -83.9266
4.4103 | -83.9255
4.4107 | -83.925
4.411 | -83.9246
4.4114 | -83.9243
4.4117 | -83.9246
4.412 | -83.9246
4.4124 | -83.9241
4.4127 | -83.9238
4.4131 | -83.9238
4.4134 | -83.9221
4.4137 | -83.9192
4.4141 | -83.9157
4.4144 | -83.9133
4.4148 | -83.9117
4.4151 | -83.9111
4.4154 | -83.9107
4.4158 | -83.9092
4.4161 | -83.9084
4.4165 | -83.9072
4.4168 | -83.905
4.4172 | -83.902
4.4175 | -83.8995
4.4178 | -83.8983
4.4182 | -83.8973
4.4185 | -83.8975
4.4189 | -83.8991
4.4192 | -83.9023
4.4195 | -83.904
4.4199 | -83.903
4.4202 | -83.9017
4.4206 | -83.9005
4.4209 | -83.8991
4.4212 | -83.8976
4.4216 | -83.8957
4.4219 | -83.8946
4.4223 | -83.8942
4.4226 | -83.8936
4.4229 | -83.893
4.4233 | -83.8928
4.4236 | -83.892
4.424 | -83.891
4.4243 | -83.8899
4.4246 | -83.8887
4.425 | -83.8876
4.4253 | -83.8863
4.4257 | -83.8852
4.426 | -83.8835
4.4263 | -83.8814
4.4267 | -83.8798
4.427 | -83.8774
4.4274 | -83.8736
4.4277 | -83.872
4.428 | -83.8707
4.4284 | -83.8703
4.4287 | -83.8689
4.4291 | -83.8663
4.4294 | -83.8644
4.4298 | -83.8633
4.4301 | -83.8618
4.4304 | -83.8612
4.4308 | -83.8597
4.4311 | -83.8587
4.4315 | -83.8584
4.4318 | -83.8579
4.4321 | -83.8573
4.4325 | -83.8562
4.4328 | -83.8548
4.4332 | -83.8538
4.4335 | -83.8529
4.4338 | -83.8515
4.4342 | -83.8495
4.4345 | -83.8473
4.4349 | -83.8447
4.4352 | -83.8426
4.4355 | -83.8414
4.4359 | -83.8395
4.4362 | -83.8362
4.4366 | -83.8337
4.4369 | -83.8326
4.4372 | -83.8314
4.4376 | -83.8299
4.4379 | -83.8287
4.4383 | -83.8272
4.4386 | -83.826
4.4389 | -83.8253
4.4393 | -83.8253
4.4396 | -83.8256
4.44 | -83.8262
4.4403 | -83.8257
4.4406 | -83.8248
4.441 | -83.824
4.4413 | -83.8247
4.4417 | -83.8269
4.442 | -83.8283
4.4423 | -83.8284
4.4427 | -83.8283
4.443 | -83.8271
4.4434 | -83.8266
4.4437 | -83.8255
4.4441 | -83.8238
4.4444 | -83.8231
4.4447 | -83.8233
4.4451 | -83.8236
4.4454 | -83.823
4.4458 | -83.8233
4.4461 | -83.8238
4.4464 | -83.8228
4.4468 | -83.8209
4.4471 | -83.8188
4.4475 | -83.8174
4.4478 | -83.8162
4.4481 | -83.8155
4.4485 | -83.8146
4.4488 | -83.8148
4.4492 | -83.8157
4.4495 | -83.8164
4.4498 | -83.8161
4.4502 | -83.815
4.4505 | -83.8131
4.4509 | -83.8119
4.4512 | -83.81
4.4515 | -83.8078
4.4519 | -83.8065
4.4522 | -83.8052
4.4526 | -83.8043
4.4529 | -83.8042
4.4532 | -83.8041
4.4536 | -83.804
4.4539 | -83.8052
4.4543 | -83.8054
4.4546 | -83.8043
4.4549 | -83.803
4.4553 | -83.8027
4.4556 | -83.8022
4.456 | -83.8015
4.4563 | -83.8003
4.4566 | -83.8
4.457 | -83.7986
4.4573 | -83.7966
4.4577 | -83.7953
4.458 | -83.794
4.4584 | -83.7934
4.4587 | -83.7929
4.459 | -83.7915
4.4594 | -83.7901
4.4597 | -83.7891
4.4601 | -83.7899
4.4604 | -83.7896
4.4607 | -83.7892
4.4611 | -83.7883
4.4614 | -83.787
4.4618 | -83.7864
4.4621 | -83.7859
4.4624 | -83.7853
4.4628 | -83.785
4.4631 | -83.7853
4.4635 | -83.785
4.4638 | -83.7848
4.4641 | -83.7845
4.4645 | -83.7839
4.4648 | -83.7834
4.4652 | -83.7834
4.4655 | -83.7835
4.4658 | -83.7831
4.4662 | -83.7824
4.4665 | -83.7818
4.4669 | -83.7822
4.4672 | -83.782
4.4675 | -83.7812
4.4679 | -83.7803
4.4682 | -83.7799
4.4686 | -83.7794
4.4689 | -83.7792
4.4692 | -83.779
4.4696 | -83.7795
4.4699 | -83.7776
4.4703 | -83.7771
4.4706 | -83.7765
4.471 | -83.7753
4.4713 | -83.7734
4.4716 | -83.7718
4.472 | -83.7717
4.4723 | -83.7721
4.4727 | -83.7728
4.473 | -83.7733
4.4733 | -83.7731
4.4737 | -83.7729
4.474 | -83.7721
4.4744 | -83.7709
4.4747 | -83.7708
4.475 | -83.7704
4.4754 | -83.7705
4.4757 | -83.7713
4.4761 | -83.7719
4.4764 | -83.7718
4.4767 | -83.7715
4.4771 | -83.7715
4.4774 | -83.7714
4.4778 | -83.7713
4.4781 | -83.7715
4.4784 | -83.7711
4.4788 | -83.7706
4.4791 | -83.7702
4.4795 | -83.7713
4.4798 | -83.7711
4.4801 | -83.771
4.4805 | -83.7704
4.4808 | -83.7706
4.4812 | -83.7709
4.4815 | -83.7711
4.4818 | -83.772
4.4822 | -83.7733
4.4825 | -83.774
4.4829 | -83.7742
4.4832 | -83.7745
4.4835 | -83.7741
4.4839 | -83.774
4.4842 | -83.774
4.4846 | -83.7738
4.4849 | -83.774
4.4852 | -83.7741
4.4856 | -83.7745
4.4859 | -83.775
4.4863 | -83.7752
4.4866 | -83.7757
4.487 | -83.7767
4.4873 | -83.7775
4.4876 | -83.7779
4.488 | -83.7786
4.4883 | -83.779
4.4887 | -83.7791
4.489 | -83.7796
4.4893 | -83.782
4.4897 | -83.7819
4.49 | -83.7825
4.4904 | -83.7831
4.4907 | -83.7836
4.491 | -83.7837
4.4914 | -83.7844
4.4917 | -83.7862
4.4921 | -83.7879
4.4924 | -83.7883
4.4927 | -83.7885
4.4931 | -83.7891
4.4934 | -83.7895
4.4938 | -83.7904
4.4941 | -83.7914
4.4944 | -83.7921
4.4948 | -83.7927
4.4951 | -83.7936
4.4955 | -83.794
4.4958 | -83.7936
4.4961 | -83.7938
4.4965 | -83.7945
4.4968 | -83.7961
4.4972 | -83.7988
4.4975 | -83.801
4.4978 | -83.8011
4.4982 | -83.8014
4.4985 | -83.8021
4.4989 | -83.803
4.4992 | -83.8049
4.4996 | -83.8059
4.4999 | -83.8076
4.5002 | -83.8102
4.5006 | -83.8116
4.5009 | -83.8121
4.5013 | -83.8135
4.5016 | -83.8143
4.5019 | -83.8151
4.5023 | -83.8169
4.5026 | -83.8183
4.503 | -83.819
4.5033 | -83.8201
4.5036 | -83.8215
4.504 | -83.8229
4.5043 | -83.8242
4.5047 | -83.8245
4.505 | -83.8246
4.5053 | -83.8261
4.5057 | -83.8281
4.506 | -83.8295
4.5064 | -83.8299
4.5067 | -83.8316
4.507 | -83.8338
4.5074 | -83.8354
4.5077 | -83.8374
4.5081 | -83.8401
4.5084 | -83.8415
4.5087 | -83.8427
4.5091 | -83.8453
4.5094 | -83.8463
4.5098 | -83.8471
4.5101 | -83.8483
4.5104 | -83.8497
4.5108 | -83.851
4.5111 | -83.8524
4.5115 | -83.8541
4.5118 | -83.8557
4.5121 | -83.8565
4.5125 | -83.8568
4.5128 | -83.8586
4.5132 | -83.8598
4.5135 | -83.8612
4.5139 | -83.8627
4.5142 | -83.8644
4.5145 | -83.8659
4.5149 | -83.867
4.5152 | -83.8684
4.5156 | -83.8701
4.5159 | -83.8715
4.5162 | -83.873
4.5166 | -83.8739
4.5169 | -83.8749
4.5173 | -83.8765
4.5176 | -83.8779
4.5179 | -83.8796
4.5183 | -83.8816
4.5186 | -83.8837
4.519 | -83.8863
4.5193 | -83.887
4.5196 | -83.8886
4.52 | -83.8901
4.5203 | -83.8915
4.5207 | -83.8929
4.521 | -83.8949
4.5213 | -83.8968
4.5217 | -83.8991
4.522 | -83.9003
4.5224 | -83.9017
4.5227 | -83.9039
4.523 | -83.9063
4.5234 | -83.9078
4.5237 | -83.9097
4.5241 | -83.9118
4.5244 | -83.9125
4.5247 | -83.9137
4.5251 | -83.9159
4.5254 | -83.9184
4.5258 | -83.9211
4.5261 | -83.9234
4.5264 | -83.9254
4.5268 | -83.9262
4.5271 | -83.928
4.5275 | -83.9307
4.5278 | -83.9341
4.5282 | -83.9372
4.5285 | -83.9406
4.5288 | -83.9419
4.5292 | -83.9424
4.5295 | -83.9444
4.5299 | -83.9465
4.5302 | -83.9492
4.5305 | -83.951
4.5309 | -83.9532
4.5312 | -83.9552
4.5316 | -83.9578
4.5319 | -83.9598
4.5322 | -83.9607
4.5326 | -83.9625
4.5329 | -83.965
4.5333 | -83.9673
4.5336 | -83.9699
4.5339 | -83.9718
4.5343 | -83.9739
4.5346 | -83.9759
4.535 | -83.9779
4.5353 | -83.9803
4.5356 | -83.9832
4.536 | -83.9858
4.5363 | -83.9875
4.5367 | -83.9896
4.537 | -83.9918
4.5373 | -83.9935
4.5377 | -83.9932
4.538 | -83.9949
4.5384 | -83.9981
4.5387 | -84.001
4.539 | -84.004
4.5394 | -84.0063
4.5397 | -84.0082
4.5401 | -84.0107
4.5404 | -84.0132
4.5408 | -84.0162
4.5411 | -84.0185
4.5414 | -84.0205
4.5418 | -84.0214
4.5421 | -84.0237
4.5425 | -84.0261
4.5428 | -84.0283
4.5431 | -84.0307
4.5435 | -84.0326
4.5438 | -84.0344
4.5442 | -84.0372
4.5445 | -84.0394
4.5448 | -84.0416
4.5452 | -84.0438
4.5455 | -84.0455
4.5459 | -84.0474
4.5462 | -84.0497
4.5465 | -84.0517
4.5469 | -84.0543
4.5472 | -84.0568
4.5476 | -84.0586
4.5479 | -84.0607
4.5482 | -84.0625
4.5486 | -84.0651
4.5489 | -84.0678
4.5493 | -84.0702
4.5496 | -84.0735
4.5499 | -84.0771
4.5503 | -84.0792
4.5506 | -84.0816
4.551 | -84.0839
4.5513 | -84.0861
4.5516 | -84.0884
4.552 | -84.0906
4.5523 | -84.0933
4.5527 | -84.0959
4.553 | -84.0982
4.5533 | -84.101
4.5537 | -84.1047
4.554 | -84.1081
4.5544 | -84.1101
4.5547 | -84.1125
4.555 | -84.1149
4.5554 | -84.117
4.5557 | -84.1195
4.5561 | -84.1221
4.5564 | -84.1248
4.5568 | -84.1277
4.5571 | -84.1302
4.5574 | -84.1324
4.5578 | -84.1353
4.5581 | -84.1382
4.5585 | -84.1401
4.5588 | -84.1441
4.5591 | -84.1478
4.5595 | -84.1507
4.5598 | -84.1543
4.5602 | -84.1573
4.5605 | -84.1597
4.5608 | -84.1626
4.5612 | -84.1643
4.5615 | -84.1666
4.5619 | -84.169
4.5622 | -84.1731
4.5625 | -84.1773
4.5629 | -84.1795
4.5632 | -84.1832
4.5636 | -84.1862
4.5639 | -84.1901
4.5642 | -84.1939
4.5646 | -84.1969
4.5649 | -84.1989
4.5653 | -84.2014
4.5656 | -84.2034
4.5659 | -84.2069
4.5663 | -84.2105
4.5666 | -84.2138
4.567 | -84.2168
4.5673 | -84.2197
4.5676 | -84.2228
4.568 | -84.2257
4.5683 | -84.2288
4.5687 | -84.2314
4.569 | -84.2344
4.5694 | -84.237
4.5697 | -84.2395
4.57 | -84.2428
4.5704 | -84.2457
4.5707 | -84.248
4.5711 | -84.2505
4.5714 | -84.2523
4.5717 | -84.256
4.5721 | -84.2586
4.5724 | -84.261
4.5728 | -84.2639
4.5731 | -84.2665
4.5734 | -84.2691
4.5738 | -84.2713
4.5741 | -84.2738
4.5745 | -84.277
4.5748 | -84.2805
4.5751 | -84.2832
4.5755 | -84.286
4.5758 | -84.2888
4.5762 | -84.2915
4.5765 | -84.2947
4.5768 | -84.2974
4.5772 | -84.2995
4.5775 | -84.302
4.5779 | -84.3049
4.5782 | -84.3081
4.5785 | -84.3103
4.5789 | -84.3135
4.5792 | -84.3163
4.5796 | -84.3199
4.5799 | -84.3247
4.5802 | -84.3296
4.5806 | -84.3323
4.5809 | -84.3343
4.5813 | -84.3373
4.5816 | -84.3403
4.5819 | -84.3432
4.5823 | -84.346
4.5826 | -84.348
4.583 | -84.3505
4.5833 | -84.3537
4.5837 | -84.3573
4.584 | -84.3604
4.5843 | -84.3632
4.5847 | -84.367
4.585 | -84.3713
4.5854 | -84.3748
4.5857 | -84.3776
4.586 | -84.3798
4.5864 | -84.3816
4.5867 | -84.384
4.5871 | -84.3867
4.5874 | -84.3898
4.5877 | -84.3903
4.5881 | -84.3951
4.5884 | -84.3995
4.5888 | -84.4025
4.5891 | -84.4049
4.5894 | -84.4073
4.5898 | -84.4096
4.5901 | -84.4122
4.5905 | -84.414
4.5908 | -84.4159
4.5911 | -84.42
4.5915 | -84.4239
4.5918 | -84.4267
4.5922 | -84.4297
4.5925 | -84.4329
4.5928 | -84.4359
4.5932 | -84.4387
4.5935 | -84.4419
4.5939 | -84.4448
4.5942 | -84.4472
4.5945 | -84.4495
4.5949 | -84.4517
4.5952 | -84.4534
4.5956 | -84.4561
4.5959 | -84.4586
4.5962 | -84.4613
4.5966 | -84.4644
4.5969 | -84.4676
4.5973 | -84.4705
4.5976 | -84.4728
4.598 | -84.4759
4.5983 | -84.4803
4.5986 | -84.4833
4.599 | -84.4869
4.5993 | -84.4911
4.5997 | -84.4941
4.6 | -84.4969
4.6003 | -84.4998
4.6007 | -84.5025
4.601 | -84.5055
4.6014 | -84.508
4.6017 | -84.5099
4.602 | -84.5123
4.6024 | -84.515
4.6027 | -84.5182
4.6031 | -84.5216
4.6034 | -84.5246
4.6037 | -84.5276
4.6041 | -84.5302
4.6044 | -84.5326
4.6048 | -84.5351
4.6051 | -84.5377
4.6054 | -84.5408
4.6058 | -84.5451
4.6061 | -84.5475
4.6065 | -84.5491
4.6068 | -84.5518
4.6071 | -84.555
4.6075 | -84.558
4.6078 | -84.5604
4.6082 | -84.5615
4.6085 | -84.5638
4.6088 | -84.5663
4.6092 | -84.5699
4.6095 | -84.5736
4.6099 | -84.5769
4.6102 | -84.5803
4.6105 | -84.583
4.6109 | -84.5857
4.6112 | -84.5882
4.6116 | -84.5906
4.6119 | -84.5932
4.6123 | -84.5953
4.6126 | -84.5977
4.6129 | -84.6003
4.6133 | -84.6032
4.6136 | -84.6056
4.614 | -84.6081
4.6143 | -84.6107
4.6146 | -84.6131
4.615 | -84.6159
4.6153 | -84.6181
4.6157 | -84.6208
4.616 | -84.6235
4.6163 | -84.625
4.6167 | -84.6277
4.617 | -84.6306
4.6174 | -84.6331
4.6177 | -84.6354
4.618 | -84.638
4.6184 | -84.6403
4.6187 | -84.6432
4.6191 | -84.6474
4.6194 | -84.6514
4.6197 | -84.6547
4.6201 | -84.6572
4.6204 | -84.6596
4.6208 | -84.6627
4.6211 | -84.6667
4.6214 | -84.6694
4.6218 | -84.6719
4.6221 | -84.675
4.6225 | -84.6764
4.6228 | -84.6789
4.6231 | -84.6818
4.6235 | -84.6848
4.6238 | -84.6878
4.6242 | -84.691
4.6245 | -84.6935
4.6249 | -84.6962
4.6252 | -84.7001
4.6255 | -84.7038
4.6259 | -84.7075
4.6262 | -84.7106
4.6266 | -84.7126
4.6269 | -84.7153
4.6272 | -84.7176
4.6276 | -84.7202
4.6279 | -84.7236
4.6283 | -84.7263
4.6286 | -84.7281
4.6289 | -84.7295
4.6293 | -84.7322
4.6296 | -84.7348
4.63 | -84.7377
4.6303 | -84.7399
4.6306 | -84.743
4.631 | -84.746
4.6313 | -84.749
4.6317 | -84.7523
4.632 | -84.7562
4.6323 | -84.7591
4.6327 | -84.7602
4.633 | -84.761
4.6334 | -84.764
4.6337 | -84.767
4.634 | -84.7693
4.6344 | -84.7725
4.6347 | -84.7755
4.6351 | -84.7786
4.6354 | -84.782
4.6357 | -84.7846
4.6361 | -84.7893
4.6364 | -84.7951
4.6368 | -84.7974
4.6371 | -84.8016
4.6374 | -84.8042
4.6378 | -84.8066
4.6381 | -84.8086
4.6385 | -84.8111
4.6388 | -84.8136
4.6392 | -84.8147
4.6395 | -84.8168
4.6398 | -84.8207
4.6402 | -84.8243
4.6405 | -84.8266
4.6409 | -84.8295
4.6412 | -84.8322
4.6415 | -84.8344
4.6419 | -84.8361
4.6422 | -84.8373
4.6426 | -84.8396
4.6429 | -84.8421
4.6432 | -84.8438
4.6436 | -84.8453
4.6439 | -84.8474
4.6443 | -84.8503
4.6446 | -84.8536
4.6449 | -84.857
4.6453 | -84.86
4.6456 | -84.8624
4.646 | -84.8648
4.6463 | -84.8672
4.6466 | -84.8688
4.647 | -84.8706
4.6473 | -84.8729
4.6477 | -84.8744
4.648 | -84.8746
4.6483 | -84.8768
4.6487 | -84.8796
4.649 | -84.8824
4.6494 | -84.8853
4.6497 | -84.8881
4.65 | -84.8896
4.6504 | -84.891
4.6507 | -84.8915
4.6511 | -84.8929
4.6514 | -84.8964
4.6517 | -84.8994
4.6521 | -84.9024
4.6524 | -84.9049
4.6528 | -84.9065
4.6531 | -84.9089
4.6535 | -84.9118
4.6538 | -84.9146
4.6541 | -84.9164
4.6545 | -84.9179
4.6548 | -84.92
4.6552 | -84.9214
4.6555 | -84.9229
4.6558 | -84.9258
4.6562 | -84.9291
4.6565 | -84.931
4.6569 | -84.9344
4.6572 | -84.9368
4.6575 | -84.9377
4.6579 | -84.9398
4.6582 | -84.9421
4.6586 | -84.9439
4.6589 | -84.9459
4.6592 | -84.9485
4.6596 | -84.9504
4.6599 | -84.9521
4.6603 | -84.9538
4.6606 | -84.9562
4.6609 | -84.9583
4.6613 | -84.9596
4.6616 | -84.961
4.662 | -84.9631
4.6623 | -84.9653
4.6626 | -84.9685
4.663 | -84.9715
4.6633 | -84.9731
4.6637 | -84.975
4.664 | -84.9768
4.6643 | -84.9787
4.6647 | -84.9805
4.665 | -84.9825
4.6654 | -84.9839
4.6657 | -84.9861
4.666 | -84.9832
4.6664 | -84.9857
4.6667 | -84.988
4.6671 | -84.9898
4.6674 | -84.9906
4.6678 | -84.9904
4.6681 | -84.99
4.6684 | -84.9902
4.6688 | -84.9923
4.6691 | -84.995
4.6695 | -84.9965
4.6698 | -84.9988
4.6701 | -85.0022
4.6705 | -85.0046
4.6708 | -85.0061
4.6712 | -85.0066
4.6715 | -85.0075
4.6718 | -85.0089
4.6722 | -85.0111
4.6725 | -85.0127
4.6729 | -85.0137
4.6732 | -85.0147
4.6735 | -85.0162
4.6739 | -85.0182
4.6742 | -85.0197
4.6746 | -85.0214
4.6749 | -85.0227
4.6752 | -85.0246
4.6756 | -85.0261
4.6759 | -85.028
4.6763 | -85.0292
4.6766 | -85.0308
4.6769 | -85.0323
4.6773 | -85.0339
4.6776 | -85.0345
4.678 | -85.0355
4.6783 | -85.0373
4.6786 | -85.0389
4.679 | -85.0399
4.6793 | -85.0404
4.6797 | -85.0422
4.68 | -85.0436
4.6803 | -85.0447
4.6807 | -85.0463
4.681 | -85.0474
4.6814 | -85.0478
4.6817 | -85.0485
4.6821 | -85.0495
4.6824 | -85.0511
4.6827 | -85.0524
4.6831 | -85.0524
4.6834 | -85.0539
4.6838 | -85.0558
4.6841 | -85.0579
4.6844 | -85.0613
4.6848 | -85.0646
4.6851 | -85.0679
4.6855 | -85.0709
4.6858 | -85.0725
4.6861 | -85.0755
4.6865 | -85.077
4.6868 | -85.0782
4.6872 | -85.0789
4.6875 | -85.08
4.6878 | -85.0812
4.6882 | -85.0824
4.6885 | -85.0842
4.6889 | -85.0856
4.6892 | -85.0859
4.6895 | -85.0874
4.6899 | -85.0889
4.6902 | -85.0894
4.6906 | -85.0911
4.6909 | -85.0929
4.6912 | -85.0938
4.6916 | -85.0955
4.6919 | -85.0967
4.6923 | -85.0978
4.6926 | -85.099
4.6929 | -85.1004
4.6933 | -85.1013
4.6936 | -85.1013
4.694 | -85.1015
4.6943 | -85.1026
4.6947 | -85.1041
4.695 | -85.1053
4.6953 | -85.1073
4.6957 | -85.1083
4.696 | -85.1101
4.6964 | -85.1106
4.6967 | -85.1103
4.697 | -85.1093
4.6974 | -85.11
4.6977 | -85.1112
4.6981 | -85.1124
4.6984 | -85.1139
4.6987 | -85.1152
4.6991 | -85.1159
4.6994 | -85.1165
4.6998 | -85.1171
4.7001 | -85.1192
4.7004 | -85.1211
4.7008 | -85.1223
4.7011 | -85.1228
4.7015 | -85.1227
4.7018 | -85.1233
4.7021 | -85.125
4.7025 | -85.1266
4.7028 | -85.1283
4.7032 | -85.1298
4.7035 | -85.1315
4.7038 | -85.1328
4.7042 | -85.1321
4.7045 | -85.1322
4.7049 | -85.1328
4.7052 | -85.1322
4.7055 | -85.1317
4.7059 | -85.1331
4.7062 | -85.1343
4.7066 | -85.1354
4.7069 | -85.1377
4.7072 | -85.139
4.7076 | -85.1402
4.7079 | -85.1413
4.7083 | -85.1423
4.7086 | -85.1439
4.709 | -85.1453
4.7093 | -85.147
4.7096 | -85.1492
4.71 | -85.1514
4.7103 | -85.1521
4.7107 | -85.1525
4.711 | -85.1528
4.7113 | -85.154
4.7117 | -85.1545
4.712 | -85.1563
4.7124 | -85.1587
4.7127 | -85.1603
4.713 | -85.1609
4.7134 | -85.1617
4.7137 | -85.1625
4.7141 | -85.1632
4.7144 | -85.1642
4.7147 | -85.1638
4.7151 | -85.1637
4.7154 | -85.1651
4.7158 | -85.1667
4.7161 | -85.1676
4.7164 | -85.1681
4.7168 | -85.169
4.7171 | -85.1706
4.7175 | -85.1724
4.7178 | -85.1739
4.7181 | -85.1731
4.7185 | -85.1731
4.7188 | -85.1726
4.7192 | -85.1711
4.7195 | -85.1709
4.7198 | -85.1714
4.7202 | -85.1717
4.7205 | -85.1722
4.7209 | -85.1737
4.7212 | -85.1754
4.7215 | -85.1773
4.7219 | -85.1789
4.7222 | -85.1799
4.7226 | -85.1814
4.7229 | -85.1851
4.7233 | -85.188
4.7236 | -85.1878
4.7239 | -85.1876
4.7243 | -85.1881
4.7246 | -85.1878
4.725 | -85.1883
4.7253 | -85.1889
4.7256 | -85.1898
4.726 | -85.1906
4.7263 | -85.1909
4.7267 | -85.1912
4.727 | -85.1926
4.7273 | -85.1935
4.7277 | -85.1934
4.728 | -85.1931
4.7284 | -85.1924
4.7287 | -85.192
4.729 | -85.1926
4.7294 | -85.1923
4.7297 | -85.1928
4.7301 | -85.1934
4.7304 | -85.1944
4.7307 | -85.195
4.7311 | -85.1966
4.7314 | -85.1988
4.7318 | -85.2015
4.7321 | -85.2039
4.7324 | -85.2053
4.7328 | -85.2062
4.7331 | -85.2063
4.7335 | -85.2064
4.7338 | -85.2067
4.7341 | -85.2076
4.7345 | -85.2091
4.7348 | -85.2069
4.7352 | -85.2091
4.7355 | -85.2106
4.7358 | -85.2098
4.7362 | -85.2098
4.7365 | -85.2103
4.7369 | -85.2111
4.7372 | -85.2108
4.7376 | -85.2112
4.7379 | -85.2124
4.7382 | -85.2127
4.7386 | -85.2126
4.7389 | -85.213
4.7393 | -85.2129
4.7396 | -85.2135
4.7399 | -85.214
4.7403 | -85.2143
4.7406 | -85.2143
4.741 | -85.2136
4.7413 | -85.2133
4.7416 | -85.2133
4.742 | -85.2133
4.7423 | -85.2133
4.7427 | -85.2128
4.743 | -85.2119
4.7433 | -85.2117
4.7437 | -85.2118
4.744 | -85.212
4.7444 | -85.2125
4.7447 | -85.2128
4.745 | -85.2132
4.7454 | -85.2131
4.7457 | -85.215
4.7461 | -85.2181
4.7464 | -85.2194
4.7467 | -85.2206
4.7471 | -85.2208
4.7474 | -85.2211
4.7478 | -85.2226
4.7481 | -85.2237
4.7484 | -85.2244
4.7488 | -85.2256
4.7491 | -85.2258
4.7495 | -85.2259
4.7498 | -85.2261
4.7501 | -85.2246
4.7505 | -85.2244
4.7508 | -85.2247
4.7512 | -85.2243
4.7515 | -85.2225
4.7519 | -85.2214
4.7522 | -85.2207
4.7525 | -85.2196
4.7529 | -85.2181
4.7532 | -85.2179
4.7536 | -85.219
4.7539 | -85.2209
4.7542 | -85.2216
4.7546 | -85.2215
4.7549 | -85.221
4.7553 | -85.2206
4.7556 | -85.2227
4.7559 | -85.2246
4.7563 | -85.2249
4.7566 | -85.2243
4.757 | -85.2231
4.7573 | -85.2219
4.7576 | -85.2211
4.758 | -85.2207
4.7583 | -85.2198
4.7587 | -85.2193
4.759 | -85.2191
4.7593 | -85.2174
4.7597 | -85.2155
4.76 | -85.2154
4.7604 | -85.215
4.7607 | -85.2137
4.761 | -85.2121
4.7614 | -85.2115
4.7617 | -85.2111
4.7621 | -85.211
4.7624 | -85.2106
4.7627 | -85.2102
4.7631 | -85.2099
4.7634 | -85.2098
4.7638 | -85.2099
4.7641 | -85.2096
4.7645 | -85.2098
4.7648 | -85.2092
4.7651 | -85.2088
4.7655 | -85.2082
4.7658 | -85.2078
4.7662 | -85.2079
4.7665 | -85.2081
4.7668 | -85.2088
4.7672 | -85.2097
4.7675 | -85.2093
4.7679 | -85.2078
4.7682 | -85.2071
4.7685 | -85.2068
4.7689 | -85.2064
4.7692 | -85.2061
4.7696 | -85.2052
4.7699 | -85.204
4.7702 | -85.2039
4.7706 | -85.2047
4.7709 | -85.2058
4.7713 | -85.2057
4.7716 | -85.206
4.7719 | -85.2063
4.7723 | -85.206
4.7726 | -85.2058
4.773 | -85.2057
4.7733 | -85.2043
4.7736 | -85.2026
4.774 | -85.2022
4.7743 | -85.2018
4.7747 | -85.2015
4.775 | -85.2011
4.7753 | -85.2008
4.7757 | -85.1995
4.776 | -85.198
4.7764 | -85.1974
4.7767 | -85.197
4.777 | -85.1974
4.7774 | -85.1978
4.7777 | -85.1988
4.7781 | -85.1996
4.7784 | -85.1991
4.7788 | -85.1991
4.7791 | -85.1981
4.7794 | -85.198
4.7798 | -85.1973
4.7801 | -85.1963
4.7805 | -85.1953
4.7808 | -85.1947
4.7811 | -85.1938
4.7815 | -85.1929
4.7818 | -85.1924
4.7822 | -85.192
4.7825 | -85.1922
4.7828 | -85.1917
4.7832 | -85.1908
4.7835 | -85.1901
4.7839 | -85.1896
4.7842 | -85.1885
4.7845 | -85.1881
4.7849 | -85.1873
4.7852 | -85.1853
4.7856 | -85.1836
4.7859 | -85.1829
4.7862 | -85.182
4.7866 | -85.1819
4.7869 | -85.1814
4.7873 | -85.1792
4.7876 | -85.1777
4.7879 | -85.1778
4.7883 | -85.1769
4.7886 | -85.175
4.789 | -85.1739
4.7893 | -85.1725
4.7896 | -85.1714
4.79 | -85.1708
4.7903 | -85.1705
4.7907 | -85.1702
4.791 | -85.1693
4.7913 | -85.1695
4.7917 | -85.1711
4.792 | -85.1707
4.7924 | -85.17
4.7927 | -85.1698
4.7931 | -85.1703
4.7934 | -85.1715
4.7937 | -85.1712
4.7941 | -85.1708
4.7944 | -85.1697
4.7948 | -85.1684
4.7951 | -85.1679
4.7954 | -85.1667
4.7958 | -85.1661
4.7961 | -85.1655
4.7965 | -85.165
4.7968 | -85.1635
4.7971 | -85.1626
4.7975 | -85.162
4.7978 | -85.1615
4.7982 | -85.1617
4.7985 | -85.1611
4.7988 | -85.1604
4.7992 | -85.1597
4.7995 | -85.1589
4.7999 | -85.1577
4.8002 | -85.1558
4.8005 | -85.1557
4.8009 | -85.1561
4.8012 | -85.1561
4.8016 | -85.1555
4.8019 | -85.1543
4.8022 | -85.1532
4.8026 | -85.1522
4.8029 | -85.1522
4.8033 | -85.1519
4.8036 | -85.1512
4.8039 | -85.1504
4.8043 | -85.1497
4.8046 | -85.1491
4.805 | -85.1486
4.8053 | -85.1481
4.8056 | -85.1481
4.806 | -85.1471
4.8063 | -85.1472
4.8067 | -85.1462
4.807 | -85.1435
4.8074 | -85.1425
4.8077 | -85.1416
4.808 | -85.1404
4.8084 | -85.1389
4.8087 | -85.1374
4.8091 | -85.1359
4.8094 | -85.1351
4.8097 | -85.1342
4.8101 | -85.1325
4.8104 | -85.1312
4.8108 | -85.1297
4.8111 | -85.1275
4.8114 | -85.1267
4.8118 | -85.1251
4.8121 | -85.1226
4.8125 | -85.1216
4.8128 | -85.1212
4.8131 | -85.1197
4.8135 | -85.1189
4.8138 | -85.1182
4.8142 | -85.1174
4.8145 | -85.1162
4.8148 | -85.1136
4.8152 | -85.1111
4.8155 | -85.1092
4.8159 | -85.1084
4.8162 | -85.1079
4.8165 | -85.1068
4.8169 | -85.1058
4.8172 | -85.1051
4.8176 | -85.1044
4.8179 | -85.1035
4.8182 | -85.1023
4.8186 | -85.1016
4.8189 | -85.1018
4.8193 | -85.1012
4.8196 | -85.1
4.82 | -85.0989
4.8203 | -85.0977
4.8206 | -85.0967
4.821 | -85.0962
4.8213 | -85.0967
4.8217 | -85.0964
4.822 | -85.0952
4.8223 | -85.0942
4.8227 | -85.094
4.823 | -85.0939
4.8234 | -85.0939
4.8237 | -85.0933
4.824 | -85.0922
4.8244 | -85.0915
4.8247 | -85.0908
4.8251 | -85.0906
4.8254 | -85.0898
4.8257 | -85.0882
4.8261 | -85.0874
4.8264 | -85.0868
4.8268 | -85.086
4.8271 | -85.0854
4.8274 | -85.0845
4.8278 | -85.0835
4.8281 | -85.0824
4.8285 | -85.0814
4.8288 | -85.0816
4.8291 | -85.082
4.8295 | -85.0818
4.8298 | -85.0808
4.8302 | -85.0794
4.8305 | -85.0781
4.8308 | -85.0775
4.8312 | -85.0764
4.8315 | -85.0753
4.8319 | -85.074
4.8322 | -85.0735
4.8325 | -85.0732
4.8329 | -85.0722
4.8332 | -85.07
4.8336 | -85.0692
4.8339 | -85.0689
4.8342 | -85.0685
4.8346 | -85.0681
4.8349 | -85.0672
4.8353 | -85.0665
4.8356 | -85.0657
4.836 | -85.0648
4.8363 | -85.0639
4.8366 | -85.0636
4.837 | -85.0638
4.8373 | -85.0642
4.8377 | -85.0633
4.838 | -85.0629
4.8383 | -85.062
4.8387 | -85.0614
4.839 | -85.061
4.8394 | -85.0597
4.8397 | -85.0589
4.84 | -85.0585
4.8404 | -85.0557
4.8407 | -85.0528
4.8411 | -85.0515
4.8414 | -85.0515
4.8417 | -85.0517
4.8421 | -85.0512
4.8424 | -85.0504
4.8428 | -85.0488
4.8431 | -85.0483
4.8434 | -85.0462
4.8438 | -85.0444
4.8441 | -85.0436
4.8445 | -85.0437
4.8448 | -85.0442
4.8451 | -85.0438
4.8455 | -85.0437
4.8458 | -85.0431
4.8462 | -85.043
4.8465 | -85.0426
4.8468 | -85.0416
4.8472 | -85.0405
4.8475 | -85.0405
4.8479 | -85.0406
4.8482 | -85.0399
4.8486 | -85.0387
4.8489 | -85.0381
4.8492 | -85.0372
4.8496 | -85.0367
4.8499 | -85.0364
4.8503 | -85.0359
4.8506 | -85.0352
4.8509 | -85.0338
4.8513 | -85.0323
4.8516 | -85.0319
4.852 | -85.032
4.8523 | -85.0322
4.8526 | -85.0315
4.853 | -85.0307
4.8533 | -85.0297
4.8537 | -85.0282
4.854 | -85.0272
4.8543 | -85.0262
4.8547 | -85.0251
4.855 | -85.0255
4.8554 | -85.0265
4.8557 | -85.0258
4.856 | -85.0248
4.8564 | -85.024
4.8567 | -85.0234
4.8571 | -85.0229
4.8574 | -85.0222
4.8577 | -85.021
4.8581 | -85.0202
4.8584 | -85.0202
4.8588 | -85.0197
4.8591 | -85.0192
4.8594 | -85.018
4.8598 | -85.017
4.8601 | -85.0164
4.8605 | -85.015
4.8608 | -85.0145
4.8611 | -85.0143
4.8615 | -85.0134
4.8618 | -85.0126
4.8622 | -85.0123
4.8625 | -85.0099
4.8629 | -85.0101
4.8632 | -85.0091
4.8635 | -85.008
4.8639 | -85.008
4.8642 | -85.0104
4.8646 | -85.0107
4.8649 | -85.0095
4.8652 | -85.0086
4.8656 | -85.0079
4.8659 | -85.0071
4.8663 | -85.0066
4.8666 | -85.0059
4.8669 | -85.005
4.8673 | -85.0044
4.8676 | -85.0036
4.868 | -85.0028
4.8683 | -85.0029
4.8686 | -85.0034
4.869 | -85.0014
4.8693 | -84.9996
4.8697 | -84.9986
4.87 | -84.998
4.8703 | -84.9977
4.8707 | -84.9973
4.871 | -84.9964
4.8714 | -84.9959
4.8717 | -84.9951
4.872 | -84.9946
4.8724 | -84.9937
4.8727 | -84.9944
4.8731 | -84.9942
4.8734 | -84.9926
4.8737 | -84.9917
4.8741 | -84.9911
4.8744 | -84.991
4.8748 | -84.9895
4.8751 | -84.9874
4.8754 | -84.9866
4.8758 | -84.987
4.8761 | -84.9855
4.8765 | -84.984
4.8768 | -84.9833
4.8772 | -84.9827
4.8775 | -84.9824
4.8778 | -84.9827
4.8782 | -84.9822
4.8785 | -84.9819
4.8789 | -84.9816
4.8792 | -84.9817
4.8795 | -84.9814
4.8799 | -84.9808
4.8802 | -84.9795
4.8806 | -84.9786
4.8809 | -84.9781
4.8812 | -84.9773
4.8816 | -84.9768
4.8819 | -84.9771
4.8823 | -84.9773
4.8826 | -84.977
4.8829 | -84.9765
4.8833 | -84.9764
4.8836 | -84.9761
4.884 | -84.9755
4.8843 | -84.9749
4.8846 | -84.9742
4.885 | -84.9743
4.8853 | -84.9742
4.8857 | -84.9736
4.886 | -84.9729
4.8863 | -84.9732
4.8867 | -84.9738
4.887 | -84.9737
4.8874 | -84.9733
4.8877 | -84.9728
4.888 | -84.9725
4.8884 | -84.9723
4.8887 | -84.972
4.8891 | -84.9712
4.8894 | -84.9724
4.8898 | -84.9726
4.8901 | -84.9724
4.8904 | -84.972
4.8908 | -84.9727
4.8911 | -84.9735
4.8915 | -84.9729
4.8918 | -84.9726
4.8921 | -84.9734
4.8925 | -84.9731
4.8928 | -84.9725
4.8932 | -84.9723
4.8935 | -84.9725
4.8938 | -84.9718
4.8942 | -84.9716
4.8945 | -84.9714
4.8949 | -84.9712
4.8952 | -84.9711
4.8955 | -84.9714
4.8959 | -84.9712
4.8962 | -84.971
4.8966 | -84.9707
4.8969 | -84.9695
4.8972 | -84.9689
4.8976 | -84.9689
4.8979 | -84.9686
4.8983 | -84.9685
4.8986 | -84.9666
4.8989 | -84.9655
4.8993 | -84.9656
4.8996 | -84.966
4.9 | -84.9661
4.9003 | -84.9669
4.9006 | -84.9657
4.901 | -84.9632
4.9013 | -84.962
4.9017 | -84.9612
4.902 | -84.9615
4.9023 | -84.9616
4.9027 | -84.9616
4.903 | -84.962
4.9034 | -84.9621
4.9037 | -84.9623
4.904 | -84.9622
4.9044 | -84.9624
4.9047 | -84.9619
4.9051 | -84.9617
4.9054 | -84.9611
4.9058 | -84.9604
4.9061 | -84.9598
4.9064 | -84.9596
4.9068 | -84.9604
4.9071 | -84.9611
4.9075 | -84.9604
4.9078 | -84.9593
4.9081 | -84.9599
4.9085 | -84.9596
4.9088 | -84.9584
4.9092 | -84.9583
4.9095 | -84.958
4.9098 | -84.9575
4.9102 | -84.9575
4.9105 | -84.9582
4.9109 | -84.959
4.9112 | -84.9597
4.9115 | -84.9598
4.9119 | -84.9608
4.9122 | -84.9605
4.9126 | -84.9603
4.9129 | -84.9601
4.9132 | -84.9599
4.9136 | -84.9596
4.9139 | -84.959
4.9143 | -84.9592
4.9146 | -84.9606
4.9149 | -84.961
4.9153 | -84.9603
4.9156 | -84.9592
4.916 | -84.9596
4.9163 | -84.9595
4.9166 | -84.9596
4.917 | -84.9596
4.9173 | -84.9592
4.9177 | -84.9599
4.918 | -84.9616
4.9184 | -84.9619
4.9187 | -84.9625
4.919 | -84.9629
4.9194 | -84.9629
4.9197 | -84.9631
4.9201 | -84.9633
4.9204 | -84.9634
4.9207 | -84.9644
4.9211 | -84.9651
4.9214 | -84.9651
4.9218 | -84.9664
4.9221 | -84.9666
4.9224 | -84.9668
4.9228 | -84.9666
4.9231 | -84.9665
4.9235 | -84.967
4.9238 | -84.9672
4.9241 | -84.9674
4.9245 | -84.9672
4.9248 | -84.9678
4.9252 | -84.969
4.9255 | -84.9698
4.9258 | -84.9707
4.9262 | -84.9703
4.9265 | -84.9701
4.9269 | -84.9704
4.9272 | -84.9708
4.9275 | -84.9713
4.9279 | -84.9702
4.9282 | -84.9689
4.9286 | -84.9696
4.9289 | -84.9706
4.9292 | -84.9707
4.9296 | -84.971
4.9299 | -84.9714
4.9303 | -84.9715
4.9306 | -84.9718
4.9309 | -84.9719
4.9313 | -84.9732
4.9316 | -84.9736
4.932 | -84.974
4.9323 | -84.9744
4.9327 | -84.9746
4.933 | -84.9756
4.9333 | -84.9758
4.9337 | -84.9758
4.934 | -84.9766
4.9344 | -84.9768
4.9347 | -84.9771
4.935 | -84.9777
4.9354 | -84.9787
4.9357 | -84.9796
4.9361 | -84.9803
4.9364 | -84.9799
4.9367 | -84.9803
4.9371 | -84.9804
4.9374 | -84.9818
4.9378 | -84.9836
4.9381 | -84.9844
4.9384 | -84.9849
4.9388 | -84.9855
4.9391 | -84.986
4.9395 | -84.987
4.9398 | -84.9891
4.9401 | -84.99
4.9405 | -84.9906
4.9408 | -84.9908
4.9412 | -84.9921
4.9415 | -84.9915
4.9418 | -84.9916
4.9422 | -84.9929
4.9425 | -84.9938
4.9429 | -84.9939
4.9432 | -84.9948
4.9435 | -84.9955
4.9439 | -84.9964
4.9442 | -84.9976
4.9446 | -84.9982
4.9449 | -84.9985
4.9452 | -84.9987
4.9456 | -84.9993
4.9459 | -85.0006
4.9463 | -85.0014
4.9466 | -85.0019
4.947 | -85.0023
4.9473 | -85.0033
4.9476 | -85.0035
4.948 | -85.004
4.9483 | -85.0047
4.9487 | -85.0047
4.949 | -85.0055
4.9493 | -85.0065
4.9497 | -85.0073
4.95 | -85.0082
4.9504 | -85.0084
4.9507 | -85.0094
4.951 | -85.0106
4.9514 | -85.0112
4.9517 | -85.0122
4.9521 | -85.0131
4.9524 | -85.0141
4.9527 | -85.0159
4.9531 | -85.0164
4.9534 | -85.0174
4.9538 | -85.0184
4.9541 | -85.0192
4.9544 | -85.0201
4.9548 | -85.0211
4.9551 | -85.0232
4.9555 | -85.0254
4.9558 | -85.0255
4.9561 | -85.0255
4.9565 | -85.0265
4.9568 | -85.0271
4.9572 | -85.0283
4.9575 | -85.0297
4.9578 | -85.031
4.9582 | -85.0319
4.9585 | -85.0325
4.9589 | -85.0338
4.9592 | -85.0361
4.9596 | -85.0371
4.9599 | -85.0379
4.9602 | -85.039
4.9606 | -85.0403
4.9609 | -85.0417
4.9613 | -85.0428
4.9616 | -85.043
4.9619 | -85.0444
4.9623 | -85.0467
4.9626 | -85.048
4.963 | -85.0489
4.9633 | -85.0498
4.9636 | -85.0503
4.964 | -85.0511
4.9643 | -85.0521
4.9647 | -85.0525
4.965 | -85.0538
4.9653 | -85.0552
4.9657 | -85.0573
4.966 | -85.0594
4.9664 | -85.0621
4.9667 | -85.0637
4.967 | -85.0649
4.9674 | -85.0654
4.9677 | -85.0655
4.9681 | -85.0664
4.9684 | -85.0672
4.9687 | -85.0685
4.9691 | -85.0703
4.9694 | -85.0714
4.9698 | -85.0728
4.9701 | -85.0744
4.9704 | -85.0754
4.9708 | -85.0776
4.9711 | -85.0786
4.9715 | -85.0795
4.9718 | -85.0808
4.9721 | -85.082
4.9725 | -85.0831
4.9728 | -85.0845
4.9732 | -85.0869
4.9735 | -85.0887
4.9739 | -85.0898
4.9742 | -85.0911
4.9745 | -85.0927
4.9749 | -85.0942
4.9752 | -85.0955
4.9756 | -85.0971
4.9759 | -85.0987
4.9762 | -85.1012
4.9766 | -85.1027
4.9769 | -85.1041
4.9773 | -85.1059
4.9776 | -85.1079
4.9779 | -85.1093
4.9783 | -85.11
4.9786 | -85.1115
4.979 | -85.1131
4.9793 | -85.1145
4.9796 | -85.1156
4.98 | -85.1174
4.9803 | -85.1188
4.9807 | -85.1209
4.981 | -85.123
4.9813 | -85.125
4.9817 | -85.127
4.982 | -85.1289
4.9824 | -85.1313
4.9827 | -85.1349
4.983 | -85.1378
4.9834 | -85.1394
4.9837 | -85.1411
4.9841 | -85.1427
4.9844 | -85.1439
4.9847 | -85.1445
4.9851 | -85.1456
4.9854 | -85.1488
4.9858 | -85.1507
4.9861 | -85.1523
4.9864 | -85.1542
4.9868 | -85.1563
4.9871 | -85.159
4.9875 | -85.1623
4.9878 | -85.1644
4.9882 | -85.1665
4.9885 | -85.1684
4.9888 | -85.1706
4.9892 | -85.173
4.9895 | -85.1747
4.9899 | -85.1763
4.9902 | -85.1783
4.9905 | -85.1811
4.9909 | -85.1833
4.9912 | -85.1856
4.9916 | -85.1876
4.9919 | -85.1895
4.9922 | -85.1916
4.9926 | -85.193
4.9929 | -85.195
4.9933 | -85.1967
4.9936 | -85.1984
4.9939 | -85.2
4.9943 | -85.2005
4.9946 | -85.2022
4.995 | -85.2046
4.9953 | -85.2059
4.9956 | -85.2067
4.996 | -85.2088
4.9963 | -85.2107
4.9967 | -85.2129
4.997 | -85.2155
4.9973 | -85.2175
4.9977 | -85.2179
4.998 | -85.2187
4.9984 | -85.221
4.9987 | -85.2227
4.999 | -85.224
4.9994 | -85.2253
4.9997 | -85.2266
5.0001 | -85.2293
5.0004 | -85.2305
5.0007 | -85.2313
5.0011 | -85.2325
5.0014 | -85.2347
5.0018 | -85.2368
5.0021 | -85.2418
5.0025 | -85.2444
5.0028 | -85.2462
5.0031 | -85.2478
5.0035 | -85.2501
5.0038 | -85.2532
5.0042 | -85.2568
5.0045 | -85.2605
5.0048 | -85.2633
5.0052 | -85.2653
5.0055 | -85.2678
5.0059 | -85.2704
5.0062 | -85.2727
5.0065 | -85.2741
5.0069 | -85.2757
5.0072 | -85.2781
5.0076 | -85.2805
5.0079 | -85.2824
5.0082 | -85.2843
5.0086 | -85.2865
5.0089 | -85.2889
5.0093 | -85.2895
5.0096 | -85.2899
5.0099 | -85.2931
5.0103 | -85.2951
5.0106 | -85.2967
5.011 | -85.2986
5.0113 | -85.3004
5.0116 | -85.302
5.012 | -85.3024
5.0123 | -85.3039
5.0127 | -85.3053
5.013 | -85.3066
5.0133 | -85.3083
5.0137 | -85.3104
5.014 | -85.312
5.0144 | -85.3142
5.0147 | -85.3165
5.015 | -85.3186
5.0154 | -85.3198
5.0157 | -85.3222
5.0161 | -85.3262
5.0164 | -85.3287
5.0168 | -85.3306
5.0171 | -85.333
5.0174 | -85.3351
5.0178 | -85.3373
5.0181 | -85.3401
5.0185 | -85.343
5.0188 | -85.3461
5.0191 | -85.3489
5.0195 | -85.3517
5.0198 | -85.354
5.0202 | -85.3573
5.0205 | -85.3599
5.0208 | -85.3613
5.0212 | -85.3627
5.0215 | -85.3644
5.0219 | -85.3661
5.0222 | -85.3676
5.0225 | -85.3699
5.0229 | -85.3707
5.0232 | -85.3725
5.0236 | -85.3747
5.0239 | -85.3768
5.0242 | -85.379
5.0246 | -85.3818
5.0249 | -85.3843
5.0253 | -85.3865
5.0256 | -85.3888
5.0259 | -85.391
5.0263 | -85.3932
5.0266 | -85.3959
5.027 | -85.3988
5.0273 | -85.4007
5.0276 | -85.4039
5.028 | -85.4069
5.0283 | -85.4097
5.0287 | -85.4119
5.029 | -85.4134
5.0293 | -85.4154
5.0297 | -85.4177
5.03 | -85.4196
5.0304 | -85.4216
5.0307 | -85.4235
5.0311 | -85.4263
5.0314 | -85.4293
5.0317 | -85.4316
5.0321 | -85.4335
5.0324 | -85.436
5.0328 | -85.4385
5.0331 | -85.4409
5.0334 | -85.4435
5.0338 | -85.446
5.0341 | -85.4488
5.0345 | -85.4524
5.0348 | -85.4557
5.0351 | -85.4602
5.0355 | -85.4625
5.0358 | -85.4649
5.0362 | -85.4673
5.0365 | -85.4689
5.0368 | -85.4711
5.0372 | -85.474
5.0375 | -85.4764
5.0379 | -85.4778
5.0382 | -85.4796
5.0385 | -85.4816
5.0389 | -85.4839
5.0392 | -85.4868
5.0396 | -85.4898
5.0399 | -85.4926
5.0402 | -85.4952
5.0406 | -85.4971
5.0409 | -85.4994
5.0413 | -85.5014
5.0416 | -85.5029
5.0419 | -85.5053
5.0423 | -85.5084
5.0426 | -85.5124
5.043 | -85.5156
5.0433 | -85.5179
5.0437 | -85.5217
5.044 | -85.5256
5.0443 | -85.5278
5.0447 | -85.5304
5.045 | -85.5324
5.0454 | -85.5348
5.0457 | -85.5375
5.046 | -85.5395
5.0464 | -85.5409
5.0467 | -85.5432
5.0471 | -85.5458
5.0474 | -85.5469
5.0477 | -85.5485
5.0481 | -85.5504
5.0484 | -85.5524
5.0488 | -85.5552
5.0491 | -85.5577
5.0494 | -85.5598
5.0498 | -85.5618
5.0501 | -85.5646
5.0505 | -85.5674
5.0508 | -85.5711
5.0511 | -85.5738
5.0515 | -85.5759
5.0518 | -85.5782
5.0522 | -85.5808
5.0525 | -85.5832
5.0528 | -85.585
5.0532 | -85.5867
5.0535 | -85.5888
5.0539 | -85.5927
5.0542 | -85.596
5.0545 | -85.5982
5.0549 | -85.6003
5.0552 | -85.6026
5.0556 | -85.6052
5.0559 | -85.6074
5.0562 | -85.6096
5.0566 | -85.6117
5.0569 | -85.6139
5.0573 | -85.6167
5.0576 | -85.6211
5.058 | -85.6244
5.0583 | -85.6272
5.0586 | -85.6312
5.059 | -85.6346
5.0593 | -85.6374
5.0597 | -85.6399
5.06 | -85.6427
5.0603 | -85.6459
5.0607 | -85.6486
5.061 | -85.6509
5.0614 | -85.6531
5.0617 | -85.6552
5.062 | -85.6574
5.0624 | -85.6596
5.0627 | -85.6619
5.0631 | -85.6647
5.0634 | -85.6681
5.0637 | -85.6703
5.0641 | -85.6725
5.0644 | -85.6757
5.0648 | -85.6786
5.0651 | -85.6807
5.0654 | -85.6833
5.0658 | -85.6862
5.0661 | -85.6887
5.0665 | -85.6911
5.0668 | -85.6929
5.0671 | -85.6952
5.0675 | -85.6979
5.0678 | -85.7009
5.0682 | -85.7043
5.0685 | -85.7078
5.0688 | -85.7104
5.0692 | -85.7131
5.0695 | -85.7152
5.0699 | -85.7186
5.0702 | -85.7209
5.0705 | -85.7224
5.0709 | -85.7245
5.0712 | -85.7261
5.0716 | -85.7278
5.0719 | -85.7298
5.0723 | -85.7323
5.0726 | -85.7338
5.0729 | -85.7354
5.0733 | -85.7376
5.0736 | -85.7401
5.074 | -85.7428
5.0743 | -85.745
5.0746 | -85.7473
5.075 | -85.7496
5.0753 | -85.7526
5.0757 | -85.7563
5.076 | -85.7573
5.0763 | -85.7595
5.0767 | -85.763
5.077 | -85.7654
5.0774 | -85.7675
5.0777 | -85.7697
5.078 | -85.7727
5.0784 | -85.7761
5.0787 | -85.7788
5.0791 | -85.781
5.0794 | -85.7827
5.0797 | -85.7853
5.0801 | -85.7882
5.0804 | -85.7906
5.0808 | -85.793
5.0811 | -85.7957
5.0814 | -85.798
5.0818 | -85.8006
5.0821 | -85.8038
5.0825 | -85.8065
5.0828 | -85.8085
5.0831 | -85.81
5.0835 | -85.8121
5.0838 | -85.8149
5.0842 | -85.8185
5.0845 | -85.8219
5.0849 | -85.8242
5.0852 | -85.8264
5.0855 | -85.8286
5.0859 | -85.831
5.0862 | -85.8334
5.0866 | -85.8366
5.0869 | -85.8397
5.0872 | -85.8413
5.0876 | -85.8439
5.0879 | -85.846
5.0883 | -85.848
5.0886 | -85.8505
5.0889 | -85.8527
5.0893 | -85.8547
5.0896 | -85.8567
5.09 | -85.8589
5.0903 | -85.8611
5.0906 | -85.8632
5.091 | -85.8646
5.0913 | -85.8669
5.0917 | -85.8698
5.092 | -85.8737
5.0923 | -85.8763
5.0927 | -85.8786
5.093 | -85.8817
5.0934 | -85.8863
5.0937 | -85.8899
5.094 | -85.8922
5.0944 | -85.8957
5.0947 | -85.9001
5.0951 | -85.9033
5.0954 | -85.9049
5.0957 | -85.9072
5.0961 | -85.9099
5.0964 | -85.9123
5.0968 | -85.9139
5.0971 | -85.9162
5.0974 | -85.9196
5.0978 | -85.9222
5.0981 | -85.9243
5.0985 | -85.9266
5.0988 | -85.9281
5.0991 | -85.9312
5.0995 | -85.935
5.0998 | -85.9377
5.1002 | -85.9405
5.1005 | -85.9424
5.1009 | -85.9448
5.1012 | -85.9475
5.1015 | -85.9496
5.1019 | -85.9522
5.1022 | -85.9557
5.1026 | -85.958
5.1029 | -85.9585
5.1032 | -85.9608
5.1036 | -85.9642
5.1039 | -85.9665
5.1043 | -85.968
5.1046 | -85.9694
5.1049 | -85.9714
5.1053 | -85.9733
5.1056 | -85.9748
5.106 | -85.9764
5.1063 | -85.9776
5.1066 | -85.9791
5.107 | -85.981
5.1073 | -85.9826
5.1077 | -85.9843
5.108 | -85.9861
5.1083 | -85.9881
5.1087 | -85.9895
5.109 | -85.9904
5.1094 | -85.9919
5.1097 | -85.9929
5.11 | -85.9939
5.1104 | -85.9957
5.1107 | -85.9982
5.1111 | -86.0016
5.1114 | -86.0037
5.1117 | -86.0051
5.1121 | -86.007
5.1124 | -86.0079
5.1128 | -86.0095
5.1131 | -86.0123
5.1135 | -86.0139
5.1138 | -86.016
5.1141 | -86.0179
5.1145 | -86.0198
5.1148 | -86.0206
5.1152 | -86.023
5.1155 | -86.0249
5.1158 | -86.0263
5.1162 | -86.027
5.1165 | -86.0292
5.1169 | -86.032
5.1172 | -86.034
5.1175 | -86.0361
5.1179 | -86.0379
5.1182 | -86.04
5.1186 | -86.0413
5.1189 | -86.0426
5.1192 | -86.0443
5.1196 | -86.047
5.1199 | -86.0499
5.1203 | -86.0518
5.1206 | -86.0535
5.1209 | -86.0558
5.1213 | -86.0579
5.1216 | -86.0587
5.122 | -86.0598
5.1223 | -86.0623
5.1226 | -86.0638
5.123 | -86.0661
5.1233 | -86.0681
5.1237 | -86.0717
5.124 | -86.0751
5.1243 | -86.0774
5.1247 | -86.0791
5.125 | -86.0812
5.1254 | -86.0831
5.1257 | -86.0841
5.126 | -86.0854
5.1264 | -86.0871
5.1267 | -86.0882
5.1271 | -86.0898
5.1274 | -86.0913
5.1278 | -86.093
5.1281 | -86.0941
5.1284 | -86.095
5.1288 | -86.0971
5.1291 | -86.0987
5.1295 | -86.1007
5.1298 | -86.1039
5.1301 | -86.1072
5.1305 | -86.1089
5.1308 | -86.1109
5.1312 | -86.113
5.1315 | -86.1145
5.1318 | -86.117
5.1322 | -86.119
5.1325 | -86.1202
5.1329 | -86.1224
5.1332 | -86.1254
5.1335 | -86.1278
5.1339 | -86.1301
5.1342 | -86.1317
5.1346 | -86.1328
5.1349 | -86.1336
5.1352 | -86.135
5.1356 | -86.1372
5.1359 | -86.1393
5.1363 | -86.14
5.1366 | -86.1408
5.1369 | -86.1431
5.1373 | -86.1463
5.1376 | -86.1483
5.138 | -86.1498
5.1383 | -86.1505
5.1386 | -86.152
5.139 | -86.1537
5.1393 | -86.1554
5.1397 | -86.1574
5.14 | -86.1597
5.1403 | -86.161
5.1407 | -86.1627
5.141 | -86.1632
5.1414 | -86.1655
5.1417 | -86.1679
5.1421 | -86.1695
5.1424 | -86.1711
5.1427 | -86.1722
5.1431 | -86.1731
5.1434 | -86.1747
5.1438 | -86.1765
5.1441 | -86.1779
5.1444 | -86.1794
5.1448 | -86.181
5.1451 | -86.1832
5.1455 | -86.186
5.1458 | -86.1877
5.1461 | -86.1885
5.1465 | -86.1894
5.1468 | -86.1915
5.1472 | -86.1939
5.1475 | -86.1956
5.1478 | -86.1971
5.1482 | -86.198
5.1485 | -86.2
5.1489 | -86.2012
5.1492 | -86.2029
5.1495 | -86.2048
5.1499 | -86.2064
5.1502 | -86.2076
5.1506 | -86.2084
5.1509 | -86.2097
5.1512 | -86.2111
5.1516 | -86.2126
5.1519 | -86.2141
5.1523 | -86.215
5.1526 | -86.2171
5.1529 | -86.2196
5.1533 | -86.221
5.1536 | -86.2227
5.154 | -86.2238
5.1543 | -86.2258
5.1547 | -86.2282
5.155 | -86.2307
5.1553 | -86.2319
5.1557 | -86.2332
5.156 | -86.2359
5.1564 | -86.2377
5.1567 | -86.2391
5.157 | -86.2418
5.1574 | -86.2425
5.1577 | -86.2444
5.1581 | -86.2452
5.1584 | -86.2481
5.1587 | -86.249
5.1591 | -86.2496
5.1594 | -86.2513
5.1598 | -86.2531
5.1601 | -86.2543
5.1604 | -86.2553
5.1608 | -86.2562
5.1611 | -86.2585
5.1615 | -86.2616
5.1618 | -86.2646
5.1621 | -86.2672
5.1625 | -86.268
5.1628 | -86.2687
5.1632 | -86.2699
5.1635 | -86.2716
5.1638 | -86.2728
5.1642 | -86.2738
5.1645 | -86.2748
5.1649 | -86.2764
5.1652 | -86.2773
5.1655 | -86.2783
5.1659 | -86.2811
5.1662 | -86.283
5.1666 | -86.2843
5.1669 | -86.2858
5.1672 | -86.2869
5.1676 | -86.2869
5.1679 | -86.2877
5.1683 | -86.2877
5.1686 | -86.2879
5.1689 | -86.2895
5.1693 | -86.2892
5.1696 | -86.2878
5.17 | -86.2866
5.1703 | -86.2866
5.1707 | -86.2873
5.171 | -86.2879
5.1713 | -86.2889
5.1717 | -86.2898
5.172 | -86.2908
5.1724 | -86.2917
5.1727 | -86.293
5.173 | -86.2957
5.1734 | -86.2973
5.1737 | -86.2976
5.1741 | -86.2998
5.1744 | -86.3017
5.1747 | -86.3032
5.1751 | -86.3042
5.1754 | -86.3056
5.1758 | -86.3068
5.1761 | -86.3072
5.1764 | -86.3082
5.1768 | -86.3089
5.1771 | -86.3088
5.1775 | -86.3103
5.1778 | -86.3128
5.1781 | -86.3151
5.1785 | -86.3142
5.1788 | -86.3153
5.1792 | -86.3168
5.1795 | -86.3187
5.1798 | -86.3195
5.1802 | -86.3203
5.1805 | -86.3212
5.1809 | -86.3228
5.1812 | -86.3232
5.1815 | -86.3232
5.1819 | -86.324
5.1822 | -86.3248
5.1826 | -86.3255
5.1829 | -86.3258
5.1833 | -86.3258
5.1836 | -86.3266
5.1839 | -86.3277
5.1843 | -86.328
5.1846 | -86.3282
5.185 | -86.3293
5.1853 | -86.3299
5.1856 | -86.3298
5.186 | -86.3307
5.1863 | -86.331
5.1867 | -86.3317
5.187 | -86.3311
5.1873 | -86.3319
5.1877 | -86.3332
5.188 | -86.3344
5.1884 | -86.3346
5.1887 | -86.3379
5.189 | -86.3417
5.1894 | -86.3429
5.1897 | -86.3432
5.1901 | -86.3435
5.1904 | -86.3436
5.1907 | -86.3442
5.1911 | -86.3452
5.1914 | -86.3459
5.1918 | -86.3463
5.1921 | -86.3485
5.1924 | -86.3506
5.1928 | -86.3512
5.1931 | -86.352
5.1935 | -86.353
5.1938 | -86.3536
5.1941 | -86.3543
5.1945 | -86.3553
5.1948 | -86.3566
5.1952 | -86.3574
5.1955 | -86.3581
5.1958 | -86.359
5.1962 | -86.3598
5.1965 | -86.3607
5.1969 | -86.3616
5.1972 | -86.3624
5.1976 | -86.3636
5.1979 | -86.3646
5.1982 | -86.3654
5.1986 | -86.3665
5.1989 | -86.3676
5.1993 | -86.3667
5.1996 | -86.3661
5.1999 | -86.3668
5.2003 | -86.3678
5.2006 | -86.3687
5.201 | -86.3691
5.2013 | -86.3701
5.2016 | -86.3714
5.202 | -86.3726
5.2023 | -86.3738
5.2027 | -86.3741
5.203 | -86.3759
5.2033 | -86.3775
5.2037 | -86.378
5.204 | -86.3786
5.2044 | -86.3791
5.2047 | -86.3793
5.205 | -86.3798
5.2054 | -86.3807
5.2057 | -86.3811
5.2061 | -86.3815
5.2064 | -86.3811
5.2067 | -86.3818
5.2071 | -86.3832
5.2074 | -86.385
5.2078 | -86.3861
5.2081 | -86.3867
5.2084 | -86.3875
5.2088 | -86.3863
5.2091 | -86.3871
5.2095 | -86.3883
5.2098 | -86.3895
5.2101 | -86.3913
5.2105 | -86.3916
5.2108 | -86.39
5.2112 | -86.389
5.2115 | -86.3897
5.2119 | -86.39
5.2122 | -86.391
5.2125 | -86.3944
5.2129 | -86.3946
5.2132 | -86.3964
5.2136 | -86.3976
5.2139 | -86.3978
5.2142 | -86.3984
5.2146 | -86.3997
5.2149 | -86.4011
5.2153 | -86.4019
5.2156 | -86.4019
5.2159 | -86.402
5.2163 | -86.4027
5.2166 | -86.4031
5.217 | -86.404
5.2173 | -86.4047
5.2176 | -86.4048
5.218 | -86.4052
5.2183 | -86.4058
5.2187 | -86.4061
5.219 | -86.4074
5.2193 | -86.4089
5.2197 | -86.4093
5.22 | -86.4098
5.2204 | -86.4099
5.2207 | -86.4099
5.221 | -86.4108
5.2214 | -86.4111
5.2217 | -86.4106
5.2221 | -86.4107
5.2224 | -86.4111
5.2227 | -86.412
5.2231 | -86.4135
5.2234 | -86.4134
5.2238 | -86.4142
5.2241 | -86.416
5.2245 | -86.4169
5.2248 | -86.4171
5.2251 | -86.4175
5.2255 | -86.4203
5.2258 | -86.4238
5.2262 | -86.4246
5.2265 | -86.4257
5.2268 | -86.4256
5.2272 | -86.427
5.2275 | -86.428
5.2279 | -86.4286
5.2282 | -86.429
5.2285 | -86.4291
5.2289 | -86.4289
5.2292 | -86.4293
5.2296 | -86.4292
5.2299 | -86.4292
5.2302 | -86.4293
5.2306 | -86.4294
5.2309 | -86.4294
5.2313 | -86.4306
5.2316 | -86.4308
5.2319 | -86.4301
5.2323 | -86.4302
5.2326 | -86.4309
5.233 | -86.4309
5.2333 | -86.431
5.2336 | -86.4311
5.234 | -86.4309
5.2343 | -86.4315
5.2347 | -86.4315
5.235 | -86.4313
5.2353 | -86.4316
5.2357 | -86.4319
5.236 | -86.4318
5.2364 | -86.4315
5.2367 | -86.4315
5.237 | -86.4318
5.2374 | -86.4323
5.2377 | -86.4329
5.2381 | -86.4332
5.2384 | -86.4338
5.2387 | -86.4337
5.2391 | -86.4327
5.2394 | -86.4319
5.2398 | -86.4319
5.2401 | -86.4306
5.2405 | -86.4292
5.2408 | -86.4293
5.2411 | -86.4294
5.2415 | -86.4297
5.2418 | -86.4299
5.2422 | -86.4303
5.2425 | -86.4306
5.2428 | -86.4311
5.2432 | -86.4312
5.2435 | -86.4317
5.2439 | -86.432
5.2442 | -86.432
5.2445 | -86.4308
5.2449 | -86.4294
5.2452 | -86.4284
5.2456 | -86.4285
5.2459 | -86.4294
5.2462 | -86.4304
5.2466 | -86.4312
5.2469 | -86.4314
5.2473 | -86.4315
5.2476 | -86.4316
5.2479 | -86.4319
5.2483 | -86.4312
5.2486 | -86.4297
5.249 | -86.4291
5.2493 | -86.4307
5.2496 | -86.4315
5.25 | -86.4317
5.2503 | -86.432
5.2507 | -86.4316
5.251 | -86.4346
5.2513 | -86.4357
5.2517 | -86.4358
5.252 | -86.4358
5.2524 | -86.4357
5.2527 | -86.4364
5.2531 | -86.4378
5.2534 | -86.441
5.2537 | -86.4428
5.2541 | -86.4439
5.2544 | -86.4448
5.2548 | -86.4451
5.2551 | -86.4449
5.2554 | -86.4447
5.2558 | -86.4451
5.2561 | -86.4454
5.2565 | -86.4453
5.2568 | -86.4452
5.2571 | -86.4452
5.2575 | -86.4456
5.2578 | -86.4459
5.2582 | -86.4452
5.2585 | -86.4455
5.2588 | -86.4457
5.2592 | -86.4457
5.2595 | -86.4472
5.2599 | -86.4473
5.2602 | -86.4471
5.2605 | -86.4479
5.2609 | -86.4493
5.2612 | -86.4493
5.2616 | -86.4498
5.2619 | -86.4512
5.2622 | -86.451
5.2626 | -86.4508
5.2629 | -86.451
5.2633 | -86.4513
5.2636 | -86.4518
5.2639 | -86.4514
5.2643 | -86.4507
5.2646 | -86.4505
5.265 | -86.4503
5.2653 | -86.4504
5.2656 | -86.4504
5.266 | -86.4502
5.2663 | -86.4511
5.2667 | -86.4519
5.267 | -86.4523
5.2674 | -86.4528
5.2677 | -86.4535
5.268 | -86.4541
5.2684 | -86.4541
5.2687 | -86.4541
5.2691 | -86.455
5.2694 | -86.456
5.2697 | -86.4561
5.2701 | -86.4576
5.2704 | -86.4578
5.2708 | -86.4573
5.2711 | -86.4584
5.2714 | -86.4594
5.2718 | -86.4592
5.2721 | -86.4591
5.2725 | -86.4592
5.2728 | -86.4588
5.2731 | -86.4591
5.2735 | -86.4595
5.2738 | -86.4602
5.2742 | -86.4597
5.2745 | -86.46
5.2748 | -86.4611
5.2752 | -86.462
5.2755 | -86.4621
5.2759 | -86.4613
5.2762 | -86.461
5.2765 | -86.4604
5.2769 | -86.46
5.2772 | -86.4607
5.2776 | -86.4608
5.2779 | -86.4606
5.2782 | -86.4606
5.2786 | -86.4607
5.2789 | -86.4614
5.2793 | -86.4617
5.2796 | -86.4612
5.2799 | -86.4603
5.2803 | -86.4606
5.2806 | -86.4616
5.281 | -86.462
5.2813 | -86.4616
5.2817 | -86.4606
5.282 | -86.4586
5.2823 | -86.4595
5.2827 | -86.462
5.283 | -86.4627
5.2834 | -86.4626
5.2837 | -86.4628
5.284 | -86.4618
5.2844 | -86.4605
5.2847 | -86.4607
5.2851 | -86.4606
5.2854 | -86.4607
5.2857 | -86.4609
5.2861 | -86.4615
5.2864 | -86.4613
5.2868 | -86.4613
5.2871 | -86.4616
5.2874 | -86.4618
5.2878 | -86.4626
5.2881 | -86.4629
5.2885 | -86.4634
5.2888 | -86.4642
5.2891 | -86.4637
5.2895 | -86.4637
5.2898 | -86.4642
5.2902 | -86.4643
5.2905 | -86.4646
5.2908 | -86.4645
5.2912 | -86.4646
5.2915 | -86.4655
5.2919 | -86.4656
5.2922 | -86.4664
5.2925 | -86.4669
5.2929 | -86.4667
5.2932 | -86.4667
5.2936 | -86.465
5.2939 | -86.4639
5.2942 | -86.4617
5.2946 | -86.4614
5.2949 | -86.4619
5.2953 | -86.4621
5.2956 | -86.4619
5.296 | -86.4624
5.2963 | -86.4627
5.2966 | -86.4617
5.297 | -86.4614
5.2973 | -86.4626
5.2977 | -86.4633
5.298 | -86.4633
5.2983 | -86.4644
5.2987 | -86.465
5.299 | -86.4649
5.2994 | -86.4649
5.2997 | -86.4649
5.3 | -86.4651
5.3004 | -86.4657
5.3007 | -86.4656
5.3011 | -86.4659
5.3014 | -86.4664
5.3017 | -86.4658
5.3021 | -86.4662
5.3024 | -86.4667
5.3028 | -86.4668
5.3031 | -86.4673
5.3034 | -86.4682
5.3038 | -86.4686
5.3041 | -86.4701
5.3045 | -86.4709
5.3048 | -86.4713
5.3051 | -86.4715
5.3055 | -86.4715
5.3058 | -86.4703
5.3062 | -86.4697
5.3065 | -86.4707
5.3068 | -86.4713
5.3072 | -86.4719
5.3075 | -86.4727
5.3079 | -86.473
5.3082 | -86.4731
5.3086 | -86.4734
5.3089 | -86.4737
5.3092 | -86.4738
5.3096 | -86.4749
5.3099 | -86.4759
5.3103 | -86.4767
5.3106 | -86.4777
5.3109 | -86.4785
5.3113 | -86.4788
5.3116 | -86.4795
5.312 | -86.4798
5.3123 | -86.4813
5.3126 | -86.482
5.313 | -86.4812
5.3133 | -86.4817
5.3137 | -86.4822
5.314 | -86.4828
5.3143 | -86.483
5.3147 | -86.484
5.315 | -86.4843
5.3154 | -86.4839
5.3157 | -86.4845
5.316 | -86.485
5.3164 | -86.4849
5.3167 | -86.4849
5.3171 | -86.4852
5.3174 | -86.4862
5.3177 | -86.4875
5.3181 | -86.4871
5.3184 | -86.4875
5.3188 | -86.4893
5.3191 | -86.4896
5.3194 | -86.4889
5.3198 | -86.4886
5.3201 | -86.4885
5.3205 | -86.4882
5.3208 | -86.4891
5.3211 | -86.4892
5.3215 | -86.4897
5.3218 | -86.4902
5.3222 | -86.4894
5.3225 | -86.4899
5.3229 | -86.4904
5.3232 | -86.4916
5.3235 | -86.4921
5.3239 | -86.4924
5.3242 | -86.4925
5.3246 | -86.4935
5.3249 | -86.4949
5.3252 | -86.4963
5.3256 | -86.4965
5.3259 | -86.4971
5.3263 | -86.4974
5.3266 | -86.4979
5.3269 | -86.4986
5.3273 | -86.4991
5.3276 | -86.4995
5.328 | -86.5003
5.3283 | -86.5011
5.3286 | -86.5012
5.329 | -86.502
5.3293 | -86.5024
5.3297 | -86.5016
5.33 | -86.5024
5.3303 | -86.5036
5.3307 | -86.5041
5.331 | -86.5044
5.3314 | -86.5048
5.3317 | -86.5055
5.332 | -86.506
5.3324 | -86.5064
5.3327 | -86.507
5.3331 | -86.5075
5.3334 | -86.5083
5.3337 | -86.5099
5.3341 | -86.5114
5.3344 | -86.5119
5.3348 | -86.5125
5.3351 | -86.5128
5.3354 | -86.5131
5.3358 | -86.5132
5.3361 | -86.5132
5.3365 | -86.5144
5.3368 | -86.5147
5.3372 | -86.5156
5.3375 | -86.516
5.3378 | -86.5162
5.3382 | -86.5163
5.3385 | -86.517
5.3389 | -86.5181
5.3392 | -86.5192
5.3395 | -86.5198
5.3399 | -86.5205
5.3402 | -86.521
5.3406 | -86.5214
5.3409 | -86.5222
5.3412 | -86.5228
5.3416 | -86.5231
5.3419 | -86.524
5.3423 | -86.5241
5.3426 | -86.5239
5.3429 | -86.5253
5.3433 | -86.5278
5.3436 | -86.5298
5.344 | -86.5311
5.3443 | -86.5318
5.3446 | -86.5329
5.345 | -86.5336
5.3453 | -86.5342
5.3457 | -86.5349
5.346 | -86.5356
5.3463 | -86.5357
5.3467 | -86.5367
5.347 | -86.5374
5.3474 | -86.5381
5.3477 | -86.5388
5.348 | -86.5398
5.3484 | -86.5412
5.3487 | -86.5413
5.3491 | -86.542
5.3494 | -86.5429
5.3497 | -86.5431
5.3501 | -86.5431
5.3504 | -86.5442
5.3508 | -86.5449
5.3511 | -86.5453
5.3515 | -86.5463
5.3518 | -86.5471
5.3521 | -86.5483
5.3525 | -86.5486
5.3528 | -86.5483
5.3532 | -86.5489
5.3535 | -86.5513
5.3538 | -86.553
5.3542 | -86.5541
5.3545 | -86.5546
5.3549 | -86.5563
5.3552 | -86.5578
5.3555 | -86.5586
5.3559 | -86.559
5.3562 | -86.5594
5.3566 | -86.5586
5.3569 | -86.5582
5.3572 | -86.5609
5.3576 | -86.5627
5.3579 | -86.5635
5.3583 | -86.5647
5.3586 | -86.5662
5.3589 | -86.5677
5.3593 | -86.5685
5.3596 | -86.5695
5.36 | -86.5706
5.3603 | -86.5721
5.3606 | -86.5742
5.361 | -86.5755
5.3613 | -86.5755
5.3617 | -86.5759
5.362 | -86.5776
5.3623 | -86.5788
5.3627 | -86.5792
5.363 | -86.5796
5.3634 | -86.5805
5.3637 | -86.5818
5.364 | -86.5834
5.3644 | -86.5841
5.3647 | -86.5844
5.3651 | -86.5851
5.3654 | -86.5855
5.3658 | -86.5852
5.3661 | -86.5853
5.3664 | -86.5864
5.3668 | -86.588
5.3671 | -86.5889
5.3675 | -86.59
5.3678 | -86.5905
5.3681 | -86.5911
5.3685 | -86.592
5.3688 | -86.5929
5.3692 | -86.5939
5.3695 | -86.597
5.3698 | -86.5998
5.3702 | -86.6006
5.3705 | -86.6008
5.3709 | -86.6011
5.3712 | -86.6023
5.3715 | -86.6011
5.3719 | -86.6015
5.3722 | -86.6027
5.3726 | -86.6052
5.3729 | -86.6068
5.3732 | -86.6084
5.3736 | -86.6096
5.3739 | -86.6113
5.3743 | -86.6126
5.3746 | -86.6137
5.3749 | -86.6152
5.3753 | -86.6167
5.3756 | -86.6179
5.376 | -86.6189
5.3763 | -86.6207
5.3766 | -86.6218
5.377 | -86.6222
5.3773 | -86.6231
5.3777 | -86.6241
5.378 | -86.6247
5.3784 | -86.6252
5.3787 | -86.6264
5.379 | -86.6277
5.3794 | -86.6282
5.3797 | -86.6292
5.3801 | -86.6298
5.3804 | -86.6306
5.3807 | -86.6318
5.3811 | -86.6334
5.3814 | -86.6335
5.3818 | -86.6342
5.3821 | -86.6363
5.3824 | -86.638
5.3828 | -86.6394
5.3831 | -86.6397
5.3835 | -86.6404
5.3838 | -86.6418
5.3841 | -86.6428
5.3845 | -86.6456
5.3848 | -86.6467
5.3852 | -86.6473
5.3855 | -86.6481
5.3858 | -86.6492
5.3862 | -86.6502
5.3865 | -86.651
5.3869 | -86.652
5.3872 | -86.6531
5.3875 | -86.6546
5.3879 | -86.6562
5.3882 | -86.6573
5.3886 | -86.6582
5.3889 | -86.6596
5.3892 | -86.6607
5.3896 | -86.6609
5.3899 | -86.6617
5.3903 | -86.6633
5.3906 | -86.6652
5.3909 | -86.6667
5.3913 | -86.6679
5.3916 | -86.6673
5.392 | -86.6673
5.3923 | -86.6686
5.3927 | -86.6703
5.393 | -86.672
5.3933 | -86.6726
5.3937 | -86.6742
5.394 | -86.6761
5.3944 | -86.6773
5.3947 | -86.6781
5.395 | -86.6795
5.3954 | -86.6809
5.3957 | -86.6822
5.3961 | -86.6833
5.3964 | -86.684
5.3967 | -86.6851
5.3971 | -86.6861
5.3974 | -86.6873
5.3978 | -86.6874
5.3981 | -86.6875
5.3984 | -86.6881
5.3988 | -86.6898
5.3991 | -86.6906
5.3995 | -86.6912
5.3998 | -86.6928
5.4001 | -86.696
5.4005 | -86.6978
5.4008 | -86.699
5.4012 | -86.7002
5.4015 | -86.701
5.4018 | -86.702
5.4022 | -86.7033
5.4025 | -86.7051
5.4029 | -86.7066
5.4032 | -86.7085
5.4035 | -86.7101
5.4039 | -86.7113
5.4042 | -86.7122
5.4046 | -86.7131
5.4049 | -86.7142
5.4052 | -86.7152
5.4056 | -86.7162
5.4059 | -86.717
5.4063 | -86.7185
5.4066 | -86.7197
5.407 | -86.7208
5.4073 | -86.7223
5.4076 | -86.7229
5.408 | -86.7247
5.4083 | -86.726
5.4087 | -86.7269
5.409 | -86.7285
5.4093 | -86.7305
5.4097 | -86.732
5.41 | -86.7334
5.4104 | -86.7356
5.4107 | -86.738
5.411 | -86.7397
5.4114 | -86.7412
5.4117 | -86.7422
5.4121 | -86.7437
5.4124 | -86.7452
5.4127 | -86.7472
5.4131 | -86.7485
5.4134 | -86.7499
5.4138 | -86.7508
5.4141 | -86.7517
5.4144 | -86.7516
5.4148 | -86.7524
5.4151 | -86.7537
5.4155 | -86.7552
5.4158 | -86.7563
5.4161 | -86.7577
5.4165 | -86.7599
5.4168 | -86.7622
5.4172 | -86.7632
5.4175 | -86.7643
5.4178 | -86.7654
5.4182 | -86.7665
5.4185 | -86.7675
5.4189 | -86.7687
5.4192 | -86.7697
5.4195 | -86.7713
5.4199 | -86.7723
5.4202 | -86.7738
5.4206 | -86.7751
5.4209 | -86.776
5.4213 | -86.7766
5.4216 | -86.7779
5.4219 | -86.7801
5.4223 | -86.7821
5.4226 | -86.7838
5.423 | -86.7851
5.4233 | -86.7863
5.4236 | -86.7877
5.424 | -86.7887
5.4243 | -86.79
5.4247 | -86.7912
5.425 | -86.7923
5.4253 | -86.7933
5.4257 | -86.794
5.426 | -86.7956
5.4264 | -86.7973
5.4267 | -86.799
5.427 | -86.8004
5.4274 | -86.8016
5.4277 | -86.8033
5.4281 | -86.8052
5.4284 | -86.8067
5.4287 | -86.8084
5.4291 | -86.8097
5.4294 | -86.8113
5.4298 | -86.8117
5.4301 | -86.8125
5.4304 | -86.8135
5.4308 | -86.8149
5.4311 | -86.8171
5.4315 | -86.8186
5.4318 | -86.8198
5.4321 | -86.8196
5.4325 | -86.8194
5.4328 | -86.8199
5.4332 | -86.8201
5.4335 | -86.8214
5.4338 | -86.8228
5.4342 | -86.8246
5.4345 | -86.8262
5.4349 | -86.8269
5.4352 | -86.8283
5.4356 | -86.8306
5.4359 | -86.8315
5.4362 | -86.8333
5.4366 | -86.8353
5.4369 | -86.8377
5.4373 | -86.8391
5.4376 | -86.8413
5.4379 | -86.843
5.4383 | -86.844
5.4386 | -86.8457
5.439 | -86.8477
5.4393 | -86.8492
5.4396 | -86.8493
5.44 | -86.8495
5.4403 | -86.8499
5.4407 | -86.8499
5.441 | -86.8525
5.4413 | -86.8539
5.4417 | -86.8544
5.442 | -86.855
5.4424 | -86.8566
5.4427 | -86.8577
5.443 | -86.8592
5.4434 | -86.8603
5.4437 | -86.8618
5.4441 | -86.8631
5.4444 | -86.8643
5.4447 | -86.8659
5.4451 | -86.868
5.4454 | -86.8695
5.4458 | -86.8703
5.4461 | -86.8715
5.4464 | -86.8725
5.4468 | -86.8737
5.4471 | -86.8749
5.4475 | -86.8762
5.4478 | -86.8774
5.4482 | -86.8781
5.4485 | -86.8786
5.4488 | -86.8798
5.4492 | -86.8814
5.4495 | -86.8826
5.4499 | -86.8837
5.4502 | -86.8842
5.4505 | -86.8846
5.4509 | -86.8837
5.4512 | -86.8844
5.4516 | -86.8857
5.4519 | -86.8868
5.4522 | -86.8878
5.4526 | -86.8884
5.4529 | -86.8897
5.4533 | -86.8916
5.4536 | -86.8921
5.4539 | -86.8927
5.4543 | -86.8937
5.4546 | -86.895
5.455 | -86.8962
5.4553 | -86.8972
5.4556 | -86.8986
5.456 | -86.9002
5.4563 | -86.9015
5.4567 | -86.9021
5.457 | -86.9026
5.4573 | -86.9043
5.4577 | -86.9055
5.458 | -86.9066
5.4584 | -86.9078
5.4587 | -86.9089
5.459 | -86.9101
5.4594 | -86.9114
5.4597 | -86.9126
5.4601 | -86.9142
5.4604 | -86.9163
5.4607 | -86.9177
5.4611 | -86.9188
5.4614 | -86.9202
5.4618 | -86.9221
5.4621 | -86.9242
5.4625 | -86.9257
5.4628 | -86.9265
5.4631 | -86.9272
5.4635 | -86.928
5.4638 | -86.9282
5.4642 | -86.929
5.4645 | -86.9298
5.4648 | -86.9307
5.4652 | -86.9307
5.4655 | -86.9322
5.4659 | -86.9346
5.4662 | -86.9365
5.4665 | -86.9375
5.4669 | -86.939
5.4672 | -86.9404
5.4676 | -86.9416
5.4679 | -86.9433
5.4682 | -86.9446
5.4686 | -86.9457
5.4689 | -86.9477
5.4693 | -86.9495
5.4696 | -86.9509
5.4699 | -86.9541
5.4703 | -86.9563
5.4706 | -86.9578
5.471 | -86.9593
5.4713 | -86.9603
5.4716 | -86.9616
5.472 | -86.9633
5.4723 | -86.9643
5.4727 | -86.9657
5.473 | -86.9674
5.4733 | -86.969
5.4737 | -86.9708
5.474 | -86.9715
5.4744 | -86.9726
5.4747 | -86.9741
5.475 | -86.9759
5.4754 | -86.9774
5.4757 | -86.9791
5.4761 | -86.9801
5.4764 | -86.9822
5.4768 | -86.985
5.4771 | -86.9865
5.4774 | -86.989
5.4778 | -86.9911
5.4781 | -86.9924
5.4785 | -86.9934
5.4788 | -86.9949
5.4791 | -86.9964
5.4795 | -86.9976
5.4798 | -86.998
5.4802 | -87.0004
5.4805 | -87.0025
5.4808 | -87.0041
5.4812 | -87.006
5.4815 | -87.0074
5.4819 | -87.0082
5.4822 | -87.0086
5.4825 | -87.0102
5.4829 | -87.013
5.4832 | -87.014
5.4836 | -87.0158
5.4839 | -87.0176
5.4842 | -87.0198
5.4846 | -87.0217
5.4849 | -87.0234
5.4853 | -87.0247
5.4856 | -87.0256
5.4859 | -87.0265
5.4863 | -87.0277
5.4866 | -87.0286
5.487 | -87.0296
5.4873 | -87.0312
5.4876 | -87.0305
5.488 | -87.0318
5.4883 | -87.033
5.4887 | -87.0341
5.489 | -87.0354
5.4893 | -87.0367
5.4897 | -87.0374
5.49 | -87.038
5.4904 | -87.0398
5.4907 | -87.0409
5.4911 | -87.0423
5.4914 | -87.0424
5.4917 | -87.0416
5.4921 | -87.0404
5.4924 | -87.0409
5.4928 | -87.042
5.4931 | -87.0428
5.4934 | -87.0438
5.4938 | -87.0445
5.4941 | -87.0456
5.4945 | -87.047
5.4948 | -87.0479
5.4951 | -87.0489
5.4955 | -87.0508
5.4958 | -87.0524
5.4962 | -87.054
5.4965 | -87.0551
5.4968 | -87.0563
5.4972 | -87.058
5.4975 | -87.0606
5.4979 | -87.0618
5.4982 | -87.0628
5.4985 | -87.0635
5.4989 | -87.0648
5.4992 | -87.0659
5.4996 | -87.0672
5.4999 | -87.0684
5.5002 | -87.0686
5.5006 | -87.0695
5.5009 | -87.0717
5.5013 | -87.0719
5.5016 | -87.074
5.5019 | -87.0757
5.5023 | -87.0765
5.5026 | -87.0769
5.503 | -87.0781
5.5033 | -87.0801
5.5037 | -87.0816
5.504 | -87.0824
5.5043 | -87.083
5.5047 | -87.084
5.505 | -87.0849
5.5054 | -87.0856
5.5057 | -87.0874
5.506 | -87.0895
5.5064 | -87.0907
5.5067 | -87.0916
5.5071 | -87.0924
5.5074 | -87.0933
5.5077 | -87.0954
5.5081 | -87.0971
5.5084 | -87.0979
5.5088 | -87.0988
5.5091 | -87.0993
5.5094 | -87.0999
5.5098 | -87.1006
5.5101 | -87.1021
5.5105 | -87.1035
5.5108 | -87.1058
5.5111 | -87.1069
5.5115 | -87.1082
5.5118 | -87.1095
5.5122 | -87.1113
5.5125 | -87.1125
5.5128 | -87.1138
5.5132 | -87.1152
5.5135 | -87.1155
5.5139 | -87.1154
5.5142 | -87.1167
5.5145 | -87.1178
5.5149 | -87.1172
5.5152 | -87.1158
5.5156 | -87.1152
5.5159 | -87.1152
5.5162 | -87.1166
5.5166 | -87.1181
5.5169 | -87.1187
5.5173 | -87.1193
5.5176 | -87.1197
5.518 | -87.1209
5.5183 | -87.1228
5.5186 | -87.1253
5.519 | -87.1263
5.5193 | -87.1267
5.5197 | -87.1273
5.52 | -87.1277
5.5203 | -87.1286
5.5207 | -87.1297
5.521 | -87.131
5.5214 | -87.1327
5.5217 | -87.1338
5.522 | -87.1346
5.5224 | -87.1349
5.5227 | -87.1365
5.5231 | -87.1371
5.5234 | -87.1371
5.5237 | -87.1402
5.5241 | -87.142
5.5244 | -87.1432
5.5248 | -87.1448
5.5251 | -87.1448
5.5254 | -87.1451
5.5258 | -87.1459
5.5261 | -87.147
5.5265 | -87.1481
5.5268 | -87.1489
5.5271 | -87.1506
5.5275 | -87.152
5.5278 | -87.1528
5.5282 | -87.1546
5.5285 | -87.1554
5.5288 | -87.1563
5.5292 | -87.1588
5.5295 | -87.16
5.5299 | -87.1612
5.5302 | -87.1635
5.5305 | -87.1652
5.5309 | -87.1668
5.5312 | -87.1672
5.5316 | -87.1665
5.5319 | -87.1665
5.5323 | -87.1674
5.5326 | -87.1685
5.5329 | -87.1701
5.5333 | -87.1712
5.5336 | -87.1722
5.534 | -87.1734
5.5343 | -87.1748
5.5346 | -87.1763
5.535 | -87.1772
5.5353 | -87.1782
5.5357 | -87.1787
5.536 | -87.1798
5.5363 | -87.1797
5.5367 | -87.1797
5.537 | -87.1797
5.5374 | -87.18
5.5377 | -87.1817
5.538 | -87.1824
5.5384 | -87.1831
5.5387 | -87.1842
5.5391 | -87.1844
5.5394 | -87.1855
5.5397 | -87.1862
5.5401 | -87.1877
5.5404 | -87.1877
5.5408 | -87.1894
5.5411 | -87.1906
5.5414 | -87.1934
5.5418 | -87.1958
5.5421 | -87.198
5.5425 | -87.2
5.5428 | -87.2009
5.5431 | -87.2018
5.5435 | -87.2021
5.5438 | -87.2024
5.5442 | -87.2028
5.5445 | -87.2032
5.5448 | -87.2031
5.5452 | -87.2042
5.5455 | -87.2052
5.5459 | -87.2064
5.5462 | -87.206
5.5466 | -87.2059
5.5469 | -87.2066
5.5472 | -87.2075
5.5476 | -87.2075
5.5479 | -87.2075
5.5483 | -87.2087
5.5486 | -87.2088
5.5489 | -87.2096
5.5493 | -87.211
5.5496 | -87.2125
5.55 | -87.2136
5.5503 | -87.2144
5.5506 | -87.2144
5.551 | -87.2151
5.5513 | -87.2178
5.5517 | -87.219
5.552 | -87.2194
5.5523 | -87.22
5.5527 | -87.2206
5.553 | -87.2214
5.5534 | -87.2218
5.5537 | -87.2225
5.554 | -87.224
5.5544 | -87.2246
5.5547 | -87.224
5.5551 | -87.224
5.5554 | -87.2247
5.5557 | -87.2259
5.5561 | -87.2268
5.5564 | -87.2275
5.5568 | -87.2282
5.5571 | -87.2294
5.5574 | -87.2306
5.5578 | -87.2308
5.5581 | -87.2309
5.5585 | -87.2325
5.5588 | -87.233
5.5591 | -87.2343
5.5595 | -87.2345
5.5598 | -87.2362
5.5602 | -87.2367
5.5605 | -87.2375
5.5609 | -87.2374
5.5612 | -87.2377
5.5615 | -87.2374
5.5619 | -87.238
5.5622 | -87.2386
5.5626 | -87.2391
5.5629 | -87.24
5.5632 | -87.2407
5.5636 | -87.2415
5.5639 | -87.2418
5.5643 | -87.2409
5.5646 | -87.2411
5.5649 | -87.2423
5.5653 | -87.2429
5.5656 | -87.2432
5.566 | -87.2433
5.5663 | -87.2431
5.5666 | -87.2432
5.567 | -87.2437
5.5673 | -87.2442
5.5677 | -87.2445
5.568 | -87.2447
5.5683 | -87.2439
5.5687 | -87.2442
5.569 | -87.245
5.5694 | -87.2456
5.5697 | -87.2464
5.57 | -87.2463
5.5704 | -87.2467
5.5707 | -87.2466
5.5711 | -87.245
5.5714 | -87.2447
5.5717 | -87.2454
5.5721 | -87.2464
5.5724 | -87.2472
5.5728 | -87.2474
5.5731 | -87.2483
5.5735 | -87.2503
5.5738 | -87.25
5.5741 | -87.2506
5.5745 | -87.2518
5.5748 | -87.2521
5.5752 | -87.2524
5.5755 | -87.2522
5.5758 | -87.2516
5.5762 | -87.2519
5.5765 | -87.2523
5.5769 | -87.2535
5.5772 | -87.2543
5.5775 | -87.2547
5.5779 | -87.2557
5.5782 | -87.2565
5.5786 | -87.2568
5.5789 | -87.2572
5.5792 | -87.2571
5.5796 | -87.2571
5.5799 | -87.2576
5.5803 | -87.2578
5.5806 | -87.2576
5.5809 | -87.2569
5.5813 | -87.2565
5.5816 | -87.2553
5.582 | -87.2549
5.5823 | -87.2539
5.5826 | -87.2541
5.583 | -87.2549
5.5833 | -87.255
5.5837 | -87.2573
5.584 | -87.2582
5.5843 | -87.2583
5.5847 | -87.2583
5.585 | -87.2585
5.5854 | -87.259
5.5857 | -87.2602
5.586 | -87.2616
5.5864 | -87.2618
5.5867 | -87.2621
5.5871 | -87.2618
5.5874 | -87.2618
5.5878 | -87.2623
5.5881 | -87.263
5.5884 | -87.2636
5.5888 | -87.264
5.5891 | -87.264
5.5895 | -87.2622
5.5898 | -87.2603
5.5901 | -87.2598
5.5905 | -87.2596
5.5908 | -87.2601
5.5912 | -87.2605
5.5915 | -87.2605
5.5918 | -87.2606
5.5922 | -87.2609
5.5925 | -87.2603
5.5929 | -87.26
5.5932 | -87.2602
5.5935 | -87.2601
5.5939 | -87.2593
5.5942 | -87.2591
5.5946 | -87.2592
5.5949 | -87.2606
5.5952 | -87.2623
5.5956 | -87.2626
5.5959 | -87.2628
5.5963 | -87.2632
5.5966 | -87.2637
5.5969 | -87.2644
5.5973 | -87.265
5.5976 | -87.2658
5.598 | -87.2656
5.5983 | -87.2647
5.5986 | -87.2644
5.599 | -87.2647
5.5993 | -87.2645
5.5997 | -87.264
5.6 | -87.2638
5.6003 | -87.2635
5.6007 | -87.263
5.601 | -87.2628
5.6014 | -87.2623
5.6017 | -87.2621
5.6021 | -87.2621
5.6024 | -87.2624
5.6027 | -87.2633
5.6031 | -87.2637
5.6034 | -87.2635
5.6038 | -87.2626
5.6041 | -87.2621
5.6044 | -87.262
5.6048 | -87.262
5.6051 | -87.2621
5.6055 | -87.2621
5.6058 | -87.2626
5.6061 | -87.2629
5.6065 | -87.2631
5.6068 | -87.2633
5.6072 | -87.2631
5.6075 | -87.2625
5.6078 | -87.2628
5.6082 | -87.2638
5.6085 | -87.2641
5.6089 | -87.264
5.6092 | -87.263
5.6095 | -87.2627
5.6099 | -87.2622
5.6102 | -87.2618
5.6106 | -87.2622
5.6109 | -87.262
5.6112 | -87.262
5.6116 | -87.2621
5.6119 | -87.262
5.6123 | -87.2618
5.6126 | -87.2615
5.6129 | -87.2615
5.6133 | -87.2623
5.6136 | -87.2626
5.614 | -87.2628
5.6143 | -87.2626
5.6146 | -87.2621
5.615 | -87.2618
5.6153 | -87.2611
5.6157 | -87.2605
5.616 | -87.2616
5.6164 | -87.2624
5.6167 | -87.2619
5.617 | -87.2639
5.6174 | -87.2653
5.6177 | -87.2655
5.6181 | -87.2651
5.6184 | -87.2642
5.6187 | -87.2633
5.6191 | -87.2624
5.6194 | -87.2614
5.6198 | -87.2603
5.6201 | -87.2588
5.6204 | -87.2588
5.6208 | -87.2592
5.6211 | -87.2589
5.6215 | -87.2589
5.6218 | -87.2601
5.6221 | -87.2613
5.6225 | -87.2623
5.6228 | -87.2622
5.6232 | -87.2617
5.6235 | -87.2613
5.6238 | -87.2611
5.6242 | -87.261
5.6245 | -87.261
5.6249 | -87.2604
5.6252 | -87.2598
5.6255 | -87.2596
5.6259 | -87.2605
5.6262 | -87.2623
5.6266 | -87.2628
5.6269 | -87.2628
5.6272 | -87.2625
5.6276 | -87.2622
5.6279 | -87.2621
5.6283 | -87.2624
5.6286 | -87.2619
5.6289 | -87.262
5.6293 | -87.2616
5.6296 | -87.2609
5.63 | -87.2607
5.6303 | -87.2604
5.6307 | -87.26
5.631 | -87.2598
5.6313 | -87.2601
5.6317 | -87.2599
5.632 | -87.2597
5.6324 | -87.2593
5.6327 | -87.259
5.633 | -87.2584
5.6334 | -87.2565
5.6337 | -87.256
5.6341 | -87.2554
5.6344 | -87.2548
5.6347 | -87.2544
5.6351 | -87.2515
5.6354 | -87.2497
5.6358 | -87.2489
5.6361 | -87.2488
5.6364 | -87.2492
5.6368 | -87.2491
5.6371 | -87.2496
5.6375 | -87.2499
5.6378 | -87.2496
5.6381 | -87.2488
5.6385 | -87.248
5.6388 | -87.2473
5.6392 | -87.2464
5.6395 | -87.2456
5.6398 | -87.2448
5.6402 | -87.2439
5.6405 | -87.2434
5.6409 | -87.2436
5.6412 | -87.2439
5.6415 | -87.2434
5.6419 | -87.2429
5.6422 | -87.2424
5.6426 | -87.2418
5.6429 | -87.2416
5.6433 | -87.2413
5.6436 | -87.2413
5.6439 | -87.2416
5.6443 | -87.2413
5.6446 | -87.2413
5.645 | -87.2417
5.6453 | -87.2419
5.6456 | -87.2415
5.646 | -87.2411
5.6463 | -87.2406
5.6467 | -87.2406
5.647 | -87.2405
5.6473 | -87.2398
5.6477 | -87.2378
5.648 | -87.235
5.6484 | -87.2348
5.6487 | -87.2323
5.649 | -87.2317
5.6494 | -87.2311
5.6497 | -87.2308
5.6501 | -87.2304
5.6504 | -87.2297
5.6507 | -87.229
5.6511 | -87.2287
5.6514 | -87.2284
5.6518 | -87.2279
5.6521 | -87.2279
5.6524 | -87.2263
5.6528 | -87.2253
5.6531 | -87.2248
5.6535 | -87.2247
5.6538 | -87.2245
5.6541 | -87.2238
5.6545 | -87.2232
5.6548 | -87.2218
5.6552 | -87.2205
5.6555 | -87.221
5.6558 | -87.2207
5.6562 | -87.2204
5.6565 | -87.2195
5.6569 | -87.2179
5.6572 | -87.2176
5.6576 | -87.2177
5.6579 | -87.2175
5.6582 | -87.217
5.6586 | -87.2168
5.6589 | -87.2163
5.6593 | -87.216
5.6596 | -87.2162
5.6599 | -87.2165
5.6603 | -87.2153
5.6606 | -87.2144
5.661 | -87.2147
5.6613 | -87.2142
5.6616 | -87.2128
5.662 | -87.214
5.6623 | -87.2142
5.6627 | -87.2137
5.663 | -87.2126
5.6633 | -87.2119
5.6637 | -87.2115
5.664 | -87.2122
5.6644 | -87.2131
5.6647 | -87.212
5.665 | -87.2118
5.6654 | -87.2112
5.6657 | -87.2107
5.6661 | -87.2103
5.6664 | -87.2096
5.6667 | -87.2087
5.6671 | -87.2079
5.6674 | -87.2073
5.6678 | -87.2066
5.6681 | -87.2061
5.6684 | -87.2047
5.6688 | -87.2034
5.6691 | -87.2027
5.6695 | -87.2023
5.6698 | -87.2024
5.6701 | -87.2023
5.6705 | -87.2007
5.6708 | -87.1998
5.6712 | -87.1988
5.6715 | -87.1978
5.6719 | -87.1975
5.6722 | -87.1976
5.6725 | -87.197
5.6729 | -87.1963
5.6732 | -87.1959
5.6736 | -87.1954
5.6739 | -87.1944
5.6742 | -87.1956
5.6746 | -87.1963
5.6749 | -87.198
5.6753 | -87.1976
5.6756 | -87.1966
5.6759 | -87.1954
5.6763 | -87.1945
5.6766 | -87.1944
5.677 | -87.1935
5.6773 | -87.1934
5.6776 | -87.1946
5.678 | -87.1947
5.6783 | -87.1946
5.6787 | -87.1944
5.679 | -87.193
5.6793 | -87.1903
5.6797 | -87.1883
5.68 | -87.1874
5.6804 | -87.186
5.6807 | -87.1852
5.681 | -87.1841
5.6814 | -87.1824
5.6817 | -87.1801
5.6821 | -87.1793
5.6824 | -87.1794
5.6827 | -87.1788
5.6831 | -87.1787
5.6834 | -87.1777
5.6838 | -87.1772
5.6841 | -87.1768
5.6844 | -87.1761
5.6848 | -87.1749
5.6851 | -87.1744
5.6855 | -87.1738
5.6858 | -87.1737
5.6862 | -87.1737
5.6865 | -87.1733
5.6868 | -87.1722
5.6872 | -87.1711
5.6875 | -87.172
5.6879 | -87.172
5.6882 | -87.1726
5.6885 | -87.1717
5.6889 | -87.171
5.6892 | -87.1702
5.6896 | -87.1698
5.6899 | -87.1693
5.6902 | -87.1683
5.6906 | -87.1672
5.6909 | -87.1666
5.6913 | -87.1661
5.6916 | -87.1657
5.6919 | -87.1651
5.6923 | -87.1644
5.6926 | -87.1638
5.693 | -87.163
5.6933 | -87.1618
5.6936 | -87.1612
5.694 | -87.1602
5.6943 | -87.1593
5.6947 | -87.1593
5.695 | -87.1591
5.6953 | -87.1583
5.6957 | -87.1571
5.696 | -87.1556
5.6964 | -87.1543
5.6967 | -87.1534
5.697 | -87.1526
5.6974 | -87.1521
5.6977 | -87.1519
5.6981 | -87.1525
5.6984 | -87.1525
5.6988 | -87.1518
5.6991 | -87.1514
5.6994 | -87.1519
5.6998 | -87.152
5.7001 | -87.1515
5.7005 | -87.1508
5.7008 | -87.1504
5.7011 | -87.1507
5.7015 | -87.1502
5.7018 | -87.1502
5.7022 | -87.1503
5.7025 | -87.1499
5.7028 | -87.1498
5.7032 | -87.1502
5.7035 | -87.1501
5.7039 | -87.15
5.7042 | -87.1491
5.7045 | -87.1493
5.7049 | -87.15
5.7052 | -87.1496
5.7056 | -87.1491
5.7059 | -87.1484
5.7062 | -87.1481
5.7066 | -87.1477
5.7069 | -87.1473
5.7073 | -87.1471
5.7076 | -87.1472
5.7079 | -87.1464
5.7083 | -87.1459
5.7086 | -87.1457
5.709 | -87.1459
5.7093 | -87.1456
5.7096 | -87.1448
5.71 | -87.1439
5.7103 | -87.1429
5.7107 | -87.1418
5.711 | -87.1395
5.7113 | -87.1374
5.7117 | -87.1364
5.712 | -87.1354
5.7124 | -87.1345
5.7127 | -87.1338
5.713 | -87.1336
5.7134 | -87.1332
5.7137 | -87.1325
5.7141 | -87.132
5.7144 | -87.1317
5.7148 | -87.1303
5.7151 | -87.1295
5.7154 | -87.1287
5.7158 | -87.1272
5.7161 | -87.1255
5.7165 | -87.1249
5.7168 | -87.1241
5.7171 | -87.1231
5.7175 | -87.1217
5.7178 | -87.1209
5.7182 | -87.1205
5.7185 | -87.12
5.7188 | -87.1186
5.7192 | -87.1165
5.7195 | -87.1155
5.7199 | -87.115
5.7202 | -87.1152
5.7205 | -87.1146
5.7209 | -87.1138
5.7212 | -87.1126
5.7216 | -87.1117
5.7219 | -87.112
5.7222 | -87.1121
5.7226 | -87.1117
5.7229 | -87.1113
5.7233 | -87.1104
5.7236 | -87.1095
5.7239 | -87.109
5.7243 | -87.1084
5.7246 | -87.1082
5.725 | -87.1076
5.7253 | -87.1072
5.7256 | -87.1067
5.726 | -87.1063
5.7263 | -87.1048
5.7267 | -87.1045
5.727 | -87.1023
5.7274 | -87.1014
5.7277 | -87.1
5.728 | -87.0993
5.7284 | -87.099
5.7287 | -87.0988
5.7291 | -87.0983
5.7294 | -87.0978
5.7297 | -87.0969
5.7301 | -87.0963
5.7304 | -87.096
5.7308 | -87.0959
5.7311 | -87.0959
5.7314 | -87.0938
5.7318 | -87.0927
5.7321 | -87.0927
5.7325 | -87.0921
5.7328 | -87.0903
5.7331 | -87.0901
5.7335 | -87.0908
5.7338 | -87.0898
5.7342 | -87.0879
5.7345 | -87.087
5.7348 | -87.0872
5.7352 | -87.0858
5.7355 | -87.0849
5.7359 | -87.0846
5.7362 | -87.0843
5.7365 | -87.0851
5.7369 | -87.0844
5.7372 | -87.0839
5.7376 | -87.0844
5.7379 | -87.0846
5.7382 | -87.0836
5.7386 | -87.0831
5.7389 | -87.0821
5.7393 | -87.0813
5.7396 | -87.081
5.7399 | -87.0807
5.7403 | -87.0795
5.7406 | -87.0799
5.741 | -87.0798
5.7413 | -87.0797
5.7417 | -87.0786
5.742 | -87.0789
5.7423 | -87.0799
5.7427 | -87.0794
5.743 | -87.079
5.7434 | -87.0787
5.7437 | -87.078
5.744 | -87.0771
5.7444 | -87.0765
5.7447 | -87.0765
5.7451 | -87.0762
5.7454 | -87.0759
5.7457 | -87.0754
5.7461 | -87.0749
5.7464 | -87.075
5.7468 | -87.0751
5.7471 | -87.0747
5.7474 | -87.0747
5.7478 | -87.0744
5.7481 | -87.0736
5.7485 | -87.0738
5.7488 | -87.0744
5.7491 | -87.0742
5.7495 | -87.0737
5.7498 | -87.0736
5.7502 | -87.0732
5.7505 | -87.073
5.7508 | -87.0727
5.7512 | -87.0719
5.7515 | -87.0723
5.7519 | -87.0732
5.7522 | -87.0735
5.7525 | -87.0726
5.7529 | -87.0725
5.7532 | -87.0723
5.7536 | -87.0715
5.7539 | -87.0706
5.7542 | -87.0698
5.7546 | -87.0689
5.7549 | -87.0686
5.7553 | -87.0684
5.7556 | -87.0683
5.756 | -87.0679
5.7563 | -87.0675
5.7566 | -87.068
5.757 | -87.0683
5.7573 | -87.0681
5.7577 | -87.0659
5.758 | -87.067
5.7583 | -87.0673
5.7587 | -87.0671
5.759 | -87.0674
5.7594 | -87.0662
5.7597 | -87.0653
5.76 | -87.0648
5.7604 | -87.0644
5.7607 | -87.0638
5.7611 | -87.0634
5.7614 | -87.063
5.7617 | -87.0627
5.7621 | -87.0618
5.7624 | -87.0608
5.7628 | -87.0595
5.7631 | -87.059
5.7634 | -87.0586
5.7638 | -87.0582
5.7641 | -87.0576
5.7645 | -87.0568
5.7648 | -87.0568
5.7651 | -87.0568
5.7655 | -87.0566
5.7658 | -87.0558
5.7662 | -87.0553
5.7665 | -87.0551
5.7668 | -87.0535
5.7672 | -87.053
5.7675 | -87.0527
5.7679 | -87.0518
5.7682 | -87.0515
5.7686 | -87.0523
5.7689 | -87.0522
5.7692 | -87.0517
5.7696 | -87.0515
5.7699 | -87.0504
5.7703 | -87.0506
5.7706 | -87.0514
5.7709 | -87.0512
5.7713 | -87.0511
5.7716 | -87.051
5.772 | -87.05
5.7723 | -87.049
5.7726 | -87.0477
5.773 | -87.0458
5.7733 | -87.0453
5.7737 | -87.0448
5.774 | -87.0445
5.7743 | -87.0439
5.7747 | -87.0427
5.775 | -87.0419
5.7754 | -87.0412
5.7757 | -87.0406
5.776 | -87.0389
5.7764 | -87.0377
5.7767 | -87.0374
5.7771 | -87.037
5.7774 | -87.0376
5.7777 | -87.0375
5.7781 | -87.0371
5.7784 | -87.0366
5.7788 | -87.0363
5.7791 | -87.0357
5.7794 | -87.0375
5.7798 | -87.039
5.7801 | -87.0382
5.7805 | -87.0373
5.7808 | -87.0368
5.7811 | -87.0364
5.7815 | -87.0361
5.7818 | -87.0351
5.7822 | -87.0337
5.7825 | -87.0334
5.7828 | -87.033
5.7832 | -87.0321
5.7835 | -87.0315
5.7839 | -87.031
5.7842 | -87.0307
5.7846 | -87.0297
5.7849 | -87.0257
5.7852 | -87.0239
5.7856 | -87.0235
5.7859 | -87.0241
5.7863 | -87.0238
5.7866 | -87.0237
5.7869 | -87.0238
5.7873 | -87.0236
5.7876 | -87.0231
5.788 | -87.0229
5.7883 | -87.0222
5.7886 | -87.0215
5.789 | -87.0213
5.7893 | -87.0211
5.7897 | -87.0199
5.79 | -87.0179
5.7903 | -87.0157
5.7907 | -87.0144
5.791 | -87.0133
5.7914 | -87.0131
5.7917 | -87.0123
5.792 | -87.0115
5.7924 | -87.011
5.7927 | -87.0105
5.7931 | -87.0108
5.7934 | -87.0112
5.7937 | -87.0104
5.7941 | -87.01
5.7944 | -87.0097
5.7948 | -87.0093
5.7951 | -87.0086
5.7954 | -87.0079
5.7958 | -87.0061
5.7961 | -87.0041
5.7965 | -87.0031
5.7968 | -87.0029
5.7972 | -87.0018
5.7975 | -87.0006
5.7978 | -86.9997
5.7982 | -86.9991
5.7985 | -86.9986
5.7989 | -86.998
5.7992 | -86.9975
5.7995 | -86.9967
5.7999 | -86.9959
5.8002 | -86.9957
5.8006 | -86.9956
5.8009 | -86.9956
5.8012 | -86.9952
5.8016 | -86.9952
5.8019 | -86.9943
5.8023 | -86.9937
5.8026 | -86.9931
5.8029 | -86.9925
5.8033 | -86.9919
5.8036 | -86.9914
5.804 | -86.9915
5.8043 | -86.9911
5.8046 | -86.9905
5.805 | -86.9901
5.8053 | -86.9896
5.8057 | -86.9897
5.806 | -86.989
5.8063 | -86.9886
5.8067 | -86.9882
5.807 | -86.9877
5.8074 | -86.9862
5.8077 | -86.9844
5.808 | -86.9828
5.8084 | -86.9821
5.8087 | -86.9816
5.8091 | -86.9809
5.8094 | -86.9806
5.8097 | -86.9797
5.8101 | -86.9781
5.8104 | -86.9775
5.8108 | -86.9769
5.8111 | -86.9761
5.8115 | -86.9753
5.8118 | -86.9748
5.8121 | -86.9741
5.8125 | -86.9729
5.8128 | -86.9719
5.8132 | -86.9713
5.8135 | -86.9708
5.8138 | -86.9701
5.8142 | -86.97
5.8145 | -86.9693
5.8149 | -86.9686
5.8152 | -86.9681
5.8155 | -86.9666
5.8159 | -86.9655
5.8162 | -86.9657
5.8166 | -86.9647
5.8169 | -86.9641
5.8172 | -86.9636
5.8176 | -86.9627
5.8179 | -86.9621
5.8183 | -86.9614
5.8186 | -86.9614
5.8189 | -86.9615
5.8193 | -86.9605
5.8196 | -86.9587
5.82 | -86.9577
5.8203 | -86.958
5.8206 | -86.9573
5.821 | -86.9566
5.8213 | -86.956
5.8217 | -86.9554
5.822 | -86.955
5.8223 | -86.9542
5.8227 | -86.9547
5.823 | -86.9547
5.8234 | -86.9544
5.8237 | -86.9545
5.824 | -86.955
5.8244 | -86.9543
5.8247 | -86.9536
5.8251 | -86.9525
5.8254 | -86.9521
5.8258 | -86.9514
5.8261 | -86.9513
5.8264 | -86.9502
5.8268 | -86.9495
5.8271 | -86.9492
5.8275 | -86.9485
5.8278 | -86.9473
5.8281 | -86.9465
5.8285 | -86.9463
5.8288 | -86.9453
5.8292 | -86.944
5.8295 | -86.9433
5.8298 | -86.9429
5.8302 | -86.9422
5.8305 | -86.941
5.8309 | -86.9405
5.8312 | -86.941
5.8315 | -86.9415
5.8319 | -86.943
5.8322 | -86.9429
5.8326 | -86.9419
5.8329 | -86.9412
5.8332 | -86.9395
5.8336 | -86.9388
5.8339 | -86.9381
5.8343 | -86.9377
5.8346 | -86.937
5.8349 | -86.9367
5.8353 | -86.936
5.8356 | -86.9352
5.836 | -86.9345
5.8363 | -86.9338
5.8366 | -86.9329
5.837 | -86.9321
5.8373 | -86.9314
5.8377 | -86.9311
5.838 | -86.9307
5.8384 | -86.93
5.8387 | -86.9289
5.839 | -86.9277
5.8394 | -86.9263
5.8397 | -86.9252
5.8401 | -86.9251
5.8404 | -86.9256
5.8407 | -86.9255
5.8411 | -86.9239
5.8414 | -86.9228
5.8418 | -86.9224
5.8421 | -86.922
5.8424 | -86.9221
5.8428 | -86.9222
5.8431 | -86.9221
5.8435 | -86.921
5.8438 | -86.9203
5.8441 | -86.9199
5.8445 | -86.9195
5.8448 | -86.9189
5.8452 | -86.9182
5.8455 | -86.9176
5.8458 | -86.9169
5.8462 | -86.9164
5.8465 | -86.9164
5.8469 | -86.9159
5.8472 | -86.9152
5.8475 | -86.9145
5.8479 | -86.914
5.8482 | -86.9129
5.8486 | -86.9119
5.8489 | -86.9111
5.8492 | -86.9102
5.8496 | -86.9088
5.8499 | -86.9086
5.8503 | -86.9076
5.8506 | -86.9066
5.8509 | -86.9065
5.8513 | -86.9062
5.8516 | -86.9052
5.852 | -86.9044
5.8523 | -86.9032
5.8526 | -86.902
5.853 | -86.9011
5.8533 | -86.9003
5.8537 | -86.8995
5.854 | -86.8978
5.8544 | -86.8968
5.8547 | -86.896
5.855 | -86.8956
5.8554 | -86.8947
5.8557 | -86.8933
5.8561 | -86.8926
5.8564 | -86.8926
5.8567 | -86.8923
5.8571 | -86.8918
5.8574 | -86.8908
5.8578 | -86.8905
5.8581 | -86.8901
5.8584 | -86.8904
5.8588 | -86.8895
5.8591 | -86.8888
5.8595 | -86.888
5.8598 | -86.8874
5.8601 | -86.8862
5.8605 | -86.8855
5.8608 | -86.885
5.8612 | -86.8839
5.8615 | -86.8829
5.8618 | -86.8826
5.8622 | -86.882
5.8625 | -86.8812
5.8629 | -86.8802
5.8632 | -86.8795
5.8635 | -86.879
5.8639 | -86.8781
5.8642 | -86.8769
5.8646 | -86.8743
5.8649 | -86.8739
5.8652 | -86.8735
5.8656 | -86.8722
5.8659 | -86.8703
5.8663 | -86.8688
5.8666 | -86.8676
5.867 | -86.8667
5.8673 | -86.866
5.8676 | -86.8655
5.868 | -86.8658
5.8683 | -86.8649
5.8687 | -86.8643
5.869 | -86.863
5.8693 | -86.8623
5.8697 | -86.8611
5.87 | -86.8603
5.8704 | -86.8592
5.8707 | -86.858
5.871 | -86.8578
5.8714 | -86.8566
5.8717 | -86.8552
5.8721 | -86.8538
5.8724 | -86.853
5.8727 | -86.852
5.8731 | -86.851
5.8734 | -86.8501
5.8738 | -86.8495
5.8741 | -86.8486
5.8744 | -86.8468
5.8748 | -86.844
5.8751 | -86.8421
5.8755 | -86.8412
5.8758 | -86.8406
5.8761 | -86.8401
5.8765 | -86.84
5.8768 | -86.8396
5.8772 | -86.8383
5.8775 | -86.8378
5.8778 | -86.8372
5.8782 | -86.8362
5.8785 | -86.8355
5.8789 | -86.8347
5.8792 | -86.8339
5.8795 | -86.8329
5.8799 | -86.8318
5.8802 | -86.8308
5.8806 | -86.8303
5.8809 | -86.8286
5.8813 | -86.8273
5.8816 | -86.8267
5.8819 | -86.8259
5.8823 | -86.8269
5.8826 | -86.826
5.883 | -86.825
5.8833 | -86.8243
5.8836 | -86.8239
5.884 | -86.8233
5.8843 | -86.8232
5.8847 | -86.8228
5.885 | -86.8233
5.8853 | -86.8224
5.8857 | -86.8209
5.886 | -86.8192
5.8864 | -86.8184
5.8867 | -86.8186
5.887 | -86.8195
5.8874 | -86.8207
5.8877 | -86.8202
5.8881 | -86.8185
5.8884 | -86.8168
5.8887 | -86.8162
5.8891 | -86.8155
5.8894 | -86.8142
5.8898 | -86.8128
5.8901 | -86.8125
5.8904 | -86.8105
5.8908 | -86.8089
5.8911 | -86.8082
5.8915 | -86.8071
5.8918 | -86.8063
5.8921 | -86.8054
5.8925 | -86.8043
5.8928 | -86.8031
5.8932 | -86.8024
5.8935 | -86.8017
5.8938 | -86.8008
5.8942 | -86.8005
5.8945 | -86.7993
5.8949 | -86.7981
5.8952 | -86.797
5.8956 | -86.7959
5.8959 | -86.7959
5.8962 | -86.7957
5.8966 | -86.7947
5.8969 | -86.7946
5.8973 | -86.7943
5.8976 | -86.7935
5.8979 | -86.7926
5.8983 | -86.7918
5.8986 | -86.7915
5.899 | -86.7909
5.8993 | -86.7902
5.8996 | -86.7896
5.9 | -86.7895
5.9003 | -86.7896
5.9007 | -86.7886
5.901 | -86.7877
5.9013 | -86.7874
5.9017 | -86.7869
5.902 | -86.7861
5.9024 | -86.7856
5.9027 | -86.7849
5.903 | -86.7843
5.9034 | -86.7835
5.9037 | -86.783
5.9041 | -86.7826
5.9044 | -86.782
5.9047 | -86.781
5.9051 | -86.7803
5.9054 | -86.7801
5.9058 | -86.78
5.9061 | -86.7796
5.9064 | -86.7788
5.9068 | -86.7781
5.9071 | -86.7774
5.9075 | -86.7772
5.9078 | -86.7767
5.9082 | -86.7761
5.9085 | -86.7755
5.9088 | -86.7746
5.9092 | -86.7743
5.9095 | -86.7738
5.9099 | -86.7726
5.9102 | -86.7726
5.9105 | -86.7729
5.9109 | -86.7727
5.9112 | -86.7721
5.9116 | -86.7729
5.9119 | -86.7743
5.9122 | -86.7751
5.9126 | -86.7758
5.9129 | -86.7752
5.9133 | -86.7744
5.9136 | -86.7741
5.9139 | -86.7745
5.9143 | -86.7745
5.9146 | -86.7745
5.915 | -86.7738
5.9153 | -86.7731
5.9156 | -86.7724
5.916 | -86.7721
5.9163 | -86.7724
5.9167 | -86.7717
5.917 | -86.7712
5.9173 | -86.7705
5.9177 | -86.7691
5.918 | -86.7684
5.9184 | -86.7681
5.9187 | -86.7674
5.919 | -86.7669
5.9194 | -86.7666
5.9197 | -86.7665
5.9201 | -86.7661
5.9204 | -86.7666
5.9207 | -86.7671
5.9211 | -86.7665
5.9214 | -86.7662
5.9218 | -86.7654
5.9221 | -86.7654
5.9225 | -86.7657
5.9228 | -86.7653
5.9231 | -86.7648
5.9235 | -86.7646
5.9238 | -86.7649
5.9242 | -86.7645
5.9245 | -86.7638
5.9248 | -86.7631
5.9252 | -86.763
5.9255 | -86.763
5.9259 | -86.7622
5.9262 | -86.7623
5.9265 | -86.7618
5.9269 | -86.7616
5.9272 | -86.7615
5.9276 | -86.7609
5.9279 | -86.7603
5.9282 | -86.76
5.9286 | -86.7595
5.9289 | -86.7591
5.9293 | -86.7592
5.9296 | -86.7604
5.9299 | -86.7609
5.9303 | -86.7607
5.9306 | -86.76
5.931 | -86.7598
5.9313 | -86.7597
5.9316 | -86.7593
5.932 | -86.7594
5.9323 | -86.7591
5.9327 | -86.7586
5.933 | -86.7584
5.9333 | -86.7591
5.9337 | -86.7583
5.934 | -86.7581
5.9344 | -86.7581
5.9347 | -86.7589
5.935 | -86.7598
5.9354 | -86.7604
5.9357 | -86.7608
5.9361 | -86.7597
5.9364 | -86.7592
5.9368 | -86.7589
5.9371 | -86.7587
5.9374 | -86.7594
5.9378 | -86.76
5.9381 | -86.76
5.9385 | -86.7601
5.9388 | -86.7605
5.9391 | -86.7609
5.9395 | -86.7607
5.9398 | -86.7607
5.9402 | -86.76
5.9405 | -86.76
5.9408 | -86.7606
5.9412 | -86.7606
5.9415 | -86.7605
5.9419 | -86.7608
5.9422 | -86.7614
5.9425 | -86.7611
5.9429 | -86.761
5.9432 | -86.7623
5.9436 | -86.7624
5.9439 | -86.7634
5.9442 | -86.7638
5.9446 | -86.7638
5.9449 | -86.7637
5.9453 | -86.7646
5.9456 | -86.7653
5.9459 | -86.7649
5.9463 | -86.7644
5.9466 | -86.7643
5.947 | -86.7645
5.9473 | -86.7643
5.9476 | -86.7645
5.948 | -86.7645
5.9483 | -86.7644
5.9487 | -86.7646
5.949 | -86.7646
5.9493 | -86.7644
5.9497 | -86.7649
5.95 | -86.7654
5.9504 | -86.7658
5.9507 | -86.7659
5.9511 | -86.766
5.9514 | -86.7661
5.9517 | -86.7665
5.9521 | -86.7668
5.9524 | -86.7668
5.9528 | -86.7673
5.9531 | -86.7687
5.9534 | -86.7697
5.9538 | -86.7707
5.9541 | -86.7719
5.9545 | -86.7731
5.9548 | -86.7738
5.9551 | -86.7744
5.9555 | -86.7748
5.9558 | -86.7755
5.9562 | -86.7755
5.9565 | -86.7754
5.9568 | -86.7757
5.9572 | -86.7761
5.9575 | -86.7768
5.9579 | -86.7776
5.9582 | -86.7781
5.9585 | -86.7785
5.9589 | -86.7789
5.9592 | -86.7798
5.9596 | -86.7804
5.9599 | -86.781
5.9602 | -86.7814
5.9606 | -86.7822
5.9609 | -86.783
5.9613 | -86.7835
5.9616 | -86.785
5.9619 | -86.786
5.9623 | -86.7866
5.9626 | -86.787
5.963 | -86.7891
5.9633 | -86.7894
5.9636 | -86.7894
5.964 | -86.7898
5.9643 | -86.7912
5.9647 | -86.7933
5.965 | -86.7948
5.9654 | -86.7954
5.9657 | -86.7963
5.966 | -86.7971
5.9664 | -86.7976
5.9667 | -86.7988
5.9671 | -86.7994
5.9674 | -86.8001
5.9677 | -86.8009
5.9681 | -86.8013
5.9684 | -86.8006
5.9688 | -86.8011
5.9691 | -86.8019
5.9694 | -86.8026
5.9698 | -86.8034
5.9701 | -86.8041
5.9705 | -86.805
5.9708 | -86.8063
5.9711 | -86.807
5.9715 | -86.808
5.9718 | -86.8107
5.9722 | -86.8121
5.9725 | -86.8129
5.9728 | -86.8142
5.9732 | -86.814
5.9735 | -86.8146
5.9739 | -86.815
5.9742 | -86.8157
5.9745 | -86.8169
5.9749 | -86.8191
5.9752 | -86.8209
5.9756 | -86.8218
5.9759 | -86.8226
5.9762 | -86.8235
5.9766 | -86.8243
5.9769 | -86.8257
5.9773 | -86.8266
5.9776 | -86.8274
5.9779 | -86.8286
5.9783 | -86.8296
5.9786 | -86.8306
5.979 | -86.8312
5.9793 | -86.8319
5.9797 | -86.8325
5.98 | -86.8334
5.9803 | -86.8348
5.9807 | -86.8359
5.981 | -86.8361
5.9814 | -86.8364
5.9817 | -86.8376
5.982 | -86.8393
5.9824 | -86.8416
5.9827 | -86.8441
5.9831 | -86.8464
5.9834 | -86.8488
5.9837 | -86.8503
5.9841 | -86.8516
5.9844 | -86.8511
5.9848 | -86.8517
5.9851 | -86.8523
5.9854 | -86.8526
5.9858 | -86.8532
5.9861 | -86.8545
5.9865 | -86.8556
5.9868 | -86.8566
5.9871 | -86.8582
5.9875 | -86.8603
5.9878 | -86.8618
5.9882 | -86.8627
5.9885 | -86.8636
5.9888 | -86.8647
5.9892 | -86.8657
5.9895 | -86.867
5.9899 | -86.8681
5.9902 | -86.8692
5.9905 | -86.8704
5.9909 | -86.8721
5.9912 | -86.8751
5.9916 | -86.8772
5.9919 | -86.8789
5.9923 | -86.8805
5.9926 | -86.8819
5.9929 | -86.8831
5.9933 | -86.8842
5.9936 | -86.8857
5.994 | -86.887
5.9943 | -86.8886
5.9946 | -86.89
5.995 | -86.892
5.9953 | -86.8938
5.9957 | -86.8952
5.996 | -86.8968
5.9963 | -86.8987
5.9967 | -86.9
5.997 | -86.9012
5.9974 | -86.9026
5.9977 | -86.9038
5.998 | -86.9047
5.9984 | -86.9055
5.9987 | -86.9069
5.9991 | -86.9085
5.9994 | -86.9102
5.9997 | -86.9116
6.0001 | -86.9122
6.0004 | -86.9136
6.0008 | -86.915
6.0011 | -86.9162
6.0014 | -86.9178
6.0018 | -86.9195
6.0021 | -86.9223
6.0025 | -86.9239
6.0028 | -86.9257
6.0031 | -86.9274
6.0035 | -86.929
6.0038 | -86.9308
6.0042 | -86.9321
6.0045 | -86.9339
6.0048 | -86.9357
6.0052 | -86.9371
6.0055 | -86.9385
6.0059 | -86.9401
6.0062 | -86.9411
6.0066 | -86.9418
6.0069 | -86.9423
6.0072 | -86.9435
6.0076 | -86.945
6.0079 | -86.9465
6.0083 | -86.9481
6.0086 | -86.9496
6.0089 | -86.9513
6.0093 | -86.953
6.0096 | -86.9548
6.01 | -86.9565
6.0103 | -86.9586
6.0106 | -86.9608
6.011 | -86.9626
6.0113 | -86.9644
6.0117 | -86.9662
6.012 | -86.9692
6.0123 | -86.9703
6.0127 | -86.972
6.013 | -86.9745
6.0134 | -86.9767
6.0137 | -86.9791
6.014 | -86.9816
6.0144 | -86.9836
6.0147 | -86.9856
6.0151 | -86.9874
6.0154 | -86.9893
6.0157 | -86.9909
6.0161 | -86.9926
6.0164 | -86.9952
6.0168 | -86.9974
6.0171 | -86.9994
6.0174 | -87.0014
6.0178 | -87.0034
6.0181 | -87.0053
6.0185 | -87.0073
6.0188 | -87.0087
6.0191 | -87.0109
6.0195 | -87.0137
6.0198 | -87.0148
6.0202 | -87.0162
6.0205 | -87.0187
6.0209 | -87.0202
6.0212 | -87.0225
6.0215 | -87.0247
6.0219 | -87.0278
6.0222 | -87.0292
6.0226 | -87.0309
6.0229 | -87.0331
6.0232 | -87.0349
6.0236 | -87.0372
6.0239 | -87.0401
6.0243 | -87.0429
6.0246 | -87.0459
6.0249 | -87.0483
6.0253 | -87.0504
6.0256 | -87.0526
6.026 | -87.0548
6.0263 | -87.0573
6.0266 | -87.0599
6.027 | -87.0621
6.0273 | -87.0644
6.0277 | -87.0667
6.028 | -87.0688
6.0283 | -87.0707
6.0287 | -87.0726
6.029 | -87.0746
6.0294 | -87.0767
6.0297 | -87.0792
6.03 | -87.0815
6.0304 | -87.0838
6.0307 | -87.086
6.0311 | -87.0888
6.0314 | -87.091
6.0317 | -87.0941
6.0321 | -87.0957
6.0324 | -87.0973
6.0328 | -87.0993
6.0331 | -87.1018
6.0335 | -87.1043
6.0338 | -87.1067
6.0341 | -87.1088
6.0345 | -87.1115
6.0348 | -87.1143
6.0352 | -87.1168
6.0355 | -87.1192
6.0358 | -87.1215
6.0362 | -87.1237
6.0365 | -87.126
6.0369 | -87.128
6.0372 | -87.1301
6.0375 | -87.1324
6.0379 | -87.1349
6.0382 | -87.1373
6.0386 | -87.1412
6.0389 | -87.1438
6.0392 | -87.1461
6.0396 | -87.149
6.0399 | -87.1518
6.0403 | -87.1567
6.0406 | -87.1606
6.0409 | -87.163
6.0413 | -87.1648
6.0416 | -87.1677
6.042 | -87.17
6.0423 | -87.1722
6.0426 | -87.1747
6.043 | -87.1772
6.0433 | -87.179
6.0437 | -87.1819
6.044 | -87.1848
6.0443 | -87.1865
6.0447 | -87.1892
6.045 | -87.1922
6.0454 | -87.1952
6.0457 | -87.1983
6.046 | -87.2014
6.0464 | -87.2041
6.0467 | -87.2063
6.0471 | -87.2083
6.0474 | -87.2113
6.0477 | -87.2143
6.0481 | -87.2171
6.0484 | -87.2201
6.0488 | -87.2225
6.0491 | -87.2244
6.0495 | -87.2257
6.0498 | -87.2269
6.0501 | -87.228
6.0505 | -87.2295
6.0508 | -87.2319
6.0512 | -87.2344
6.0515 | -87.237
6.0518 | -87.2396
6.0522 | -87.2417
6.0525 | -87.2443
6.0529 | -87.2468
6.0532 | -87.249
6.0535 | -87.2513
6.0539 | -87.2552
6.0542 | -87.2603
6.0546 | -87.2635
6.0549 | -87.2673
6.0552 | -87.2707
6.0556 | -87.2735
6.0559 | -87.276
6.0563 | -87.2785
6.0566 | -87.2815
6.0569 | -87.2842
6.0573 | -87.2862
6.0576 | -87.2885
6.058 | -87.2908
6.0583 | -87.2929
6.0586 | -87.2955
6.059 | -87.2991
6.0593 | -87.3022
6.0597 | -87.3049
6.06 | -87.3076
6.0603 | -87.3098
6.0607 | -87.312
6.061 | -87.3137
6.0614 | -87.3159
6.0617 | -87.3182
6.0621 | -87.3202
6.0624 | -87.3225
6.0627 | -87.3252
6.0631 | -87.3278
6.0634 | -87.3311
6.0638 | -87.3347
6.0641 | -87.3373
6.0644 | -87.34
6.0648 | -87.3421
6.0651 | -87.344
6.0655 | -87.3462
6.0658 | -87.3489
6.0661 | -87.3514
6.0665 | -87.3538
6.0668 | -87.3564
6.0672 | -87.3589
6.0675 | -87.3619
6.0678 | -87.365
6.0682 | -87.3674
6.0685 | -87.3701
6.0689 | -87.3738
6.0692 | -87.3752
6.0695 | -87.3777
6.0699 | -87.3802
6.0702 | -87.3829
6.0706 | -87.3854
6.0709 | -87.3883
6.0712 | -87.3907
6.0716 | -87.3931
6.0719 | -87.3956
6.0723 | -87.3986
6.0726 | -87.4014
6.0729 | -87.4041
6.0733 | -87.4063
6.0736 | -87.4081
6.074 | -87.4106
6.0743 | -87.4137
6.0746 | -87.4165
6.075 | -87.419
6.0753 | -87.4214
6.0757 | -87.4237
6.076 | -87.4265
6.0764 | -87.43
6.0767 | -87.4332
6.077 | -87.4362
6.0774 | -87.4385
6.0777 | -87.4406
6.0781 | -87.443
6.0784 | -87.4462
6.0787 | -87.4488
6.0791 | -87.4518
6.0794 | -87.4549
6.0798 | -87.4575
6.0801 | -87.4599
6.0804 | -87.4625
6.0808 | -87.4648
6.0811 | -87.4676
6.0815 | -87.4704
6.0818 | -87.4732
6.0821 | -87.4761
6.0825 | -87.4789
6.0828 | -87.4816
6.0832 | -87.4842
6.0835 | -87.4861
6.0838 | -87.4888
6.0842 | -87.4918
6.0845 | -87.4945
6.0849 | -87.4971
6.0852 | -87.4999
6.0855 | -87.5019
6.0859 | -87.5042
6.0862 | -87.5064
6.0866 | -87.5083
6.0869 | -87.5107
6.0872 | -87.5132
6.0876 | -87.5159
6.0879 | -87.5202
6.0883 | -87.524
6.0886 | -87.5271
6.0889 | -87.5304
6.0893 | -87.5332
6.0896 | -87.5363
6.09 | -87.5396
6.0903 | -87.5434
6.0907 | -87.5459
6.091 | -87.5493
6.0913 | -87.5517
6.0917 | -87.554
6.092 | -87.5557
6.0924 | -87.5582
6.0927 | -87.5605
6.093 | -87.5626
6.0934 | -87.5651
6.0937 | -87.568
6.0941 | -87.5705
6.0944 | -87.5734
6.0947 | -87.5758
6.0951 | -87.5779
6.0954 | -87.5803
6.0958 | -87.5832
6.0961 | -87.5861
6.0964 | -87.5886
6.0968 | -87.5908
6.0971 | -87.5932
6.0975 | -87.5957
6.0978 | -87.5982
6.0981 | -87.6004
6.0985 | -87.6028
6.0988 | -87.6053
6.0992 | -87.608
6.0995 | -87.6108
6.0998 | -87.6137
6.1002 | -87.6163
6.1005 | -87.6181
6.1009 | -87.6199
6.1012 | -87.6221
6.1015 | -87.6244
6.1019 | -87.6264
6.1022 | -87.6281
6.1026 | -87.6302
6.1029 | -87.6331
6.1033 | -87.6355
6.1036 | -87.6383
6.1039 | -87.6408
6.1043 | -87.6424
6.1046 | -87.6455
6.105 | -87.6483
6.1053 | -87.6513
6.1056 | -87.6545
6.106 | -87.6577
6.1063 | -87.6604
6.1067 | -87.6631
6.107 | -87.6658
6.1073 | -87.6684
6.1077 | -87.6709
6.108 | -87.6732
6.1084 | -87.6749
6.1087 | -87.6763
6.109 | -87.6787
6.1094 | -87.681
6.1097 | -87.6832
6.1101 | -87.6855
6.1104 | -87.6877
6.1107 | -87.6903
6.1111 | -87.6934
6.1114 | -87.6958
6.1118 | -87.6973
6.1121 | -87.6989
6.1124 | -87.7009
6.1128 | -87.7035
6.1131 | -87.7057
6.1135 | -87.7078
6.1138 | -87.7099
6.1141 | -87.7125
6.1145 | -87.7152
6.1148 | -87.7177
6.1152 | -87.7198
6.1155 | -87.7221
6.1158 | -87.7243
6.1162 | -87.7267
6.1165 | -87.729
6.1169 | -87.731
6.1172 | -87.7331
6.1175 | -87.7353
6.1179 | -87.7374
6.1182 | -87.74
6.1186 | -87.7427
6.1189 | -87.7442
6.1193 | -87.7455
6.1196 | -87.7475
6.1199 | -87.7497
6.1203 | -87.7511
6.1206 | -87.7535
6.121 | -87.7562
6.1213 | -87.7587
6.1216 | -87.7608
6.122 | -87.7649
6.1223 | -87.7687
6.1227 | -87.7716
6.123 | -87.7748
6.1233 | -87.7771
6.1237 | -87.7802
6.124 | -87.7834
6.1244 | -87.7855
6.1247 | -87.7878
6.125 | -87.7901
6.1254 | -87.7923
6.1257 | -87.7952
6.1261 | -87.7978
6.1264 | -87.7998
6.1267 | -87.8029
6.1271 | -87.8063
6.1274 | -87.8086
6.1278 | -87.8107
6.1281 | -87.8134
6.1284 | -87.8163
6.1288 | -87.8195
6.1291 | -87.822
6.1295 | -87.8238
6.1298 | -87.8248
6.1301 | -87.8269
6.1305 | -87.83
6.1308 | -87.8339
6.1312 | -87.8366
6.1315 | -87.8391
6.1319 | -87.8407
6.1322 | -87.842
6.1325 | -87.8441
6.1329 | -87.8469
6.1332 | -87.8494
6.1336 | -87.8513
6.1339 | -87.8524
6.1342 | -87.8535
6.1346 | -87.8552
6.1349 | -87.8576
6.1353 | -87.8599
6.1356 | -87.8615
6.1359 | -87.8634
6.1363 | -87.8653
6.1366 | -87.8672
6.137 | -87.8693
6.1373 | -87.8721
6.1376 | -87.8739
6.138 | -87.8767
6.1383 | -87.8795
6.1387 | -87.8829
6.139 | -87.8857
6.1393 | -87.8877
6.1397 | -87.8885
6.14 | -87.8913
6.1404 | -87.8935
6.1407 | -87.8942
6.141 | -87.8951
6.1414 | -87.8968
6.1417 | -87.8985
6.1421 | -87.9005
6.1424 | -87.902
6.1427 | -87.9035
6.1431 | -87.9051
6.1434 | -87.9073
6.1438 | -87.9105
6.1441 | -87.913
6.1444 | -87.9146
6.1448 | -87.9163
6.1451 | -87.9191
6.1455 | -87.9203
6.1458 | -87.9215
6.1462 | -87.9237
6.1465 | -87.9265
6.1468 | -87.9287
6.1472 | -87.9303
6.1475 | -87.9322
6.1479 | -87.9344
6.1482 | -87.9358
6.1485 | -87.937
6.1489 | -87.9381
6.1492 | -87.94
6.1496 | -87.9414
6.1499 | -87.9427
6.1502 | -87.9428
6.1506 | -87.9443
6.1509 | -87.9459
6.1513 | -87.9471
6.1516 | -87.9486
6.1519 | -87.9506
6.1523 | -87.9528
6.1526 | -87.9546
6.153 | -87.9559
6.1533 | -87.958
6.1536 | -87.9583
6.154 | -87.9598
6.1543 | -87.9613
6.1547 | -87.9641
6.155 | -87.9661
6.1553 | -87.9667
6.1557 | -87.9675
6.156 | -87.9687
6.1564 | -87.9698
6.1567 | -87.9714
6.157 | -87.9729
6.1574 | -87.9741
6.1577 | -87.9758
6.1581 | -87.9771
6.1584 | -87.9775
6.1587 | -87.9788
6.1591 | -87.9804
6.1594 | -87.9816
6.1598 | -87.983
6.1601 | -87.9846
6.1605 | -87.986
6.1608 | -87.9866
6.1611 | -87.9866
6.1615 | -87.9874
6.1618 | -87.9884
6.1622 | -87.9894
6.1625 | -87.9899
6.1628 | -87.9914
6.1632 | -87.9919
6.1635 | -87.9899
6.1639 | -87.989
6.1642 | -87.9889
6.1645 | -87.99
6.1649 | -87.9919
6.1652 | -87.9929
6.1656 | -87.9939
6.1659 | -87.9954
6.1662 | -87.9975
6.1666 | -87.9984
6.1669 | -87.9993
6.1673 | -88.001
6.1676 | -88.0029
6.1679 | -88.0044
6.1683 | -88.0056
6.1686 | -88.0066
6.169 | -88.0075
6.1693 | -88.0084
6.1696 | -88.0092
6.17 | -88.0098
6.1703 | -88.0108
6.1707 | -88.0119
6.171 | -88.013
6.1713 | -88.0134
6.1717 | -88.0145
6.172 | -88.0178
6.1724 | -88.0196
6.1727 | -88.0206
6.173 | -88.0214
6.1734 | -88.0222
6.1737 | -88.0242
6.1741 | -88.0257
6.1744 | -88.0272
6.1748 | -88.0288
6.1751 | -88.0301
6.1754 | -88.0312
6.1758 | -88.0321
6.1761 | -88.0329
6.1765 | -88.0345
6.1768 | -88.036
6.1771 | -88.0375
6.1775 | -88.0382
6.1778 | -88.0393
6.1782 | -88.0406
6.1785 | -88.0421
6.1788 | -88.0435
6.1792 | -88.0443
6.1795 | -88.0444
6.1799 | -88.0444
6.1802 | -88.0453
6.1805 | -88.0467
6.1809 | -88.0484
6.1812 | -88.0486
6.1816 | -88.0498
6.1819 | -88.052
6.1822 | -88.0534
6.1826 | -88.054
6.1829 | -88.0558
6.1833 | -88.057
6.1836 | -88.0579
6.1839 | -88.0594
6.1843 | -88.0605
6.1846 | -88.0613
6.185 | -88.0629
6.1853 | -88.0645
6.1856 | -88.0652
6.186 | -88.0658
6.1863 | -88.0666
6.1867 | -88.0675
6.187 | -88.0689
6.1874 | -88.0705
6.1877 | -88.073
6.188 | -88.0755
6.1884 | -88.0765
6.1887 | -88.0775
6.1891 | -88.0785
6.1894 | -88.0793
6.1897 | -88.0804
6.1901 | -88.082
6.1904 | -88.0825
6.1908 | -88.0838
6.1911 | -88.0858
6.1914 | -88.0881
6.1918 | -88.0887
6.1921 | -88.0884
6.1925 | -88.0891
6.1928 | -88.0898
6.1931 | -88.0905
6.1935 | -88.0909
6.1938 | -88.0891
6.1942 | -88.0892
6.1945 | -88.0898
6.1948 | -88.0908
6.1952 | -88.0922
6.1955 | -88.0929
6.1959 | -88.0944
6.1962 | -88.0957
6.1965 | -88.0966
6.1969 | -88.0974
6.1972 | -88.0977
6.1976 | -88.0985
6.1979 | -88.0995
6.1982 | -88.0999
6.1986 | -88.1007
6.1989 | -88.1016
6.1993 | -88.1022
6.1996 | -88.1029
6.1999 | -88.1033
6.2003 | -88.1038
6.2006 | -88.1056
6.201 | -88.11
6.2013 | -88.1129
6.2017 | -88.1138
6.202 | -88.1145
6.2023 | -88.1155
6.2027 | -88.116
6.203 | -88.1171
6.2034 | -88.117
6.2037 | -88.1176
6.204 | -88.1185
6.2044 | -88.1193
6.2047 | -88.1202
6.2051 | -88.1211
6.2054 | -88.1216
6.2057 | -88.1219
6.2061 | -88.1228
6.2064 | -88.1252
6.2068 | -88.1278
6.2071 | -88.1294
6.2074 | -88.1302
6.2078 | -88.1306
6.2081 | -88.1316
6.2085 | -88.1323
6.2088 | -88.1326
6.2091 | -88.1327
6.2095 | -88.1331
6.2098 | -88.1337
6.2102 | -88.1344
6.2105 | -88.1351
6.2108 | -88.1356
6.2112 | -88.1359
6.2115 | -88.1367
6.2119 | -88.1369
6.2122 | -88.1374
6.2125 | -88.1385
6.2129 | -88.1405
6.2132 | -88.1428
6.2136 | -88.1437
6.2139 | -88.1438
6.2142 | -88.1446
6.2146 | -88.1461
6.2149 | -88.1472
6.2153 | -88.1478
6.2156 | -88.1481
6.216 | -88.1486
6.2163 | -88.1484
6.2166 | -88.1493
6.217 | -88.15
6.2173 | -88.1503
6.2177 | -88.1507
6.218 | -88.151
6.2183 | -88.1517
6.2187 | -88.1518
6.219 | -88.1521
6.2194 | -88.1528
6.2197 | -88.1533
6.22 | -88.1544
6.2204 | -88.1556
6.2207 | -88.1565
6.2211 | -88.1573
6.2214 | -88.1584
6.2217 | -88.159
6.2221 | -88.159
6.2224 | -88.159
6.2228 | -88.159
6.2231 | -88.1592
6.2234 | -88.1583
6.2238 | -88.1578
6.2241 | -88.1578
6.2245 | -88.1584
6.2248 | -88.1593
6.2251 | -88.161
6.2255 | -88.163
6.2258 | -88.1649
6.2262 | -88.1659
6.2265 | -88.1657
6.2268 | -88.1654
6.2272 | -88.1654
6.2275 | -88.1669
6.2279 | -88.1686
6.2282 | -88.1696
6.2285 | -88.17
6.2289 | -88.1704
6.2292 | -88.1713
6.2296 | -88.1717
6.2299 | -88.1723
6.2303 | -88.172
6.2306 | -88.1722
6.2309 | -88.1724
6.2313 | -88.1727
6.2316 | -88.1729
6.232 | -88.1734
6.2323 | -88.1735
6.2326 | -88.1741
6.233 | -88.1747
6.2333 | -88.175
6.2337 | -88.1748
6.234 | -88.1743
6.2343 | -88.1743
6.2347 | -88.174
6.235 | -88.1736
6.2354 | -88.1735
6.2357 | -88.1737
6.236 | -88.1745
6.2364 | -88.1745
6.2367 | -88.1737
6.2371 | -88.1737
6.2374 | -88.174
6.2377 | -88.1741
6.2381 | -88.1746
6.2384 | -88.1763
6.2388 | -88.1772
6.2391 | -88.1775
6.2394 | -88.1784
6.2398 | -88.1786
6.2401 | -88.1787
6.2405 | -88.1785
6.2408 | -88.1762
6.2411 | -88.1753
6.2415 | -88.1757
6.2418 | -88.177
6.2422 | -88.1775
6.2425 | -88.1776
6.2428 | -88.1773
6.2432 | -88.1773
6.2435 | -88.1776
6.2439 | -88.1784
6.2442 | -88.1784
6.2446 | -88.1784
6.2449 | -88.1787
6.2452 | -88.1788
6.2456 | -88.1789
6.2459 | -88.1783
6.2463 | -88.1782
6.2466 | -88.1784
6.2469 | -88.1789
6.2473 | -88.1794
6.2476 | -88.1792
6.248 | -88.1775
6.2483 | -88.1773
6.2486 | -88.1787
6.249 | -88.1791
6.2493 | -88.1777
6.2497 | -88.1757
6.25 | -88.1758
6.2503 | -88.1765
6.2507 | -88.176
6.251 | -88.1762
6.2514 | -88.1754
6.2517 | -88.1746
6.252 | -88.1749
6.2524 | -88.1759
6.2527 | -88.1777
6.2531 | -88.1774
6.2534 | -88.1769
6.2537 | -88.1769
6.2541 | -88.1763
6.2544 | -88.1758
6.2548 | -88.1758
6.2551 | -88.1752
6.2554 | -88.1746
6.2558 | -88.1743
6.2561 | -88.173
6.2565 | -88.173
6.2568 | -88.1732
6.2572 | -88.1726
6.2575 | -88.172
6.2578 | -88.1711
6.2582 | -88.1706
6.2585 | -88.1708
6.2589 | -88.1715
6.2592 | -88.1718
6.2595 | -88.1724
6.2599 | -88.1729
6.2602 | -88.1726
6.2606 | -88.1724
6.2609 | -88.1727
6.2612 | -88.174
6.2616 | -88.1748
6.2619 | -88.1749
6.2623 | -88.175
6.2626 | -88.175
6.2629 | -88.1752
6.2633 | -88.1745
6.2636 | -88.1742
6.264 | -88.174
6.2643 | -88.1738
6.2646 | -88.1733
6.265 | -88.1737
6.2653 | -88.174
6.2657 | -88.1735
6.266 | -88.1724
6.2663 | -88.1724
6.2667 | -88.1721
6.267 | -88.1716
6.2674 | -88.1716
6.2677 | -88.1705
6.268 | -88.1694
6.2684 | -88.1686
6.2687 | -88.1676
6.2691 | -88.1671
6.2694 | -88.1668
6.2697 | -88.165
6.2701 | -88.1638
6.2704 | -88.1631
6.2708 | -88.1605
6.2711 | -88.1583
6.2715 | -88.1569
6.2718 | -88.1553
6.2721 | -88.1532
6.2725 | -88.1517
6.2728 | -88.1504
6.2732 | -88.1484
6.2735 | -88.1456
6.2738 | -88.145
6.2742 | -88.1447
6.2745 | -88.1441
6.2749 | -88.1446
6.2752 | -88.1445
6.2755 | -88.1437
6.2759 | -88.143
6.2762 | -88.1422
6.2766 | -88.1418
6.2769 | -88.1421
6.2772 | -88.1413
6.2776 | -88.1414
6.2779 | -88.141
6.2783 | -88.1415
6.2786 | -88.1423
6.2789 | -88.1424
6.2793 | -88.1419
6.2796 | -88.1401
6.28 | -88.1385
6.2803 | -88.1382
6.2806 | -88.1376
6.281 | -88.1366
6.2813 | -88.1364
6.2817 | -88.1368
6.282 | -88.1363
6.2823 | -88.1339
6.2827 | -88.1319
6.283 | -88.1313
6.2834 | -88.1307
6.2837 | -88.1304
6.284 | -88.13
6.2844 | -88.1293
6.2847 | -88.1276
6.2851 | -88.1264
6.2854 | -88.1257
6.2858 | -88.1259
6.2861 | -88.1268
6.2864 | -88.1265
6.2868 | -88.1263
6.2871 | -88.1256
6.2875 | -88.1249
6.2878 | -88.1247
6.2881 | -88.1253
6.2885 | -88.1253
6.2888 | -88.1247
6.2892 | -88.1241
6.2895 | -88.1227
6.2898 | -88.1205
6.2902 | -88.1189
6.2905 | -88.1183
6.2909 | -88.1168
6.2912 | -88.1167
6.2915 | -88.1162
6.2919 | -88.1161
6.2922 | -88.1157
6.2926 | -88.1144
6.2929 | -88.1121
6.2932 | -88.1094
6.2936 | -88.1086
6.2939 | -88.1083
6.2943 | -88.1079
6.2946 | -88.1073
6.2949 | -88.106
6.2953 | -88.1054
6.2956 | -88.1051
6.296 | -88.1046
6.2963 | -88.1037
6.2966 | -88.1032
6.297 | -88.1025
6.2973 | -88.1031
6.2977 | -88.1046
6.298 | -88.1044
6.2983 | -88.103
6.2987 | -88.102
6.299 | -88.1012
6.2994 | -88.0996
6.2997 | -88.0987
6.3001 | -88.0984
6.3004 | -88.0968
6.3007 | -88.0956
6.3011 | -88.0944
6.3014 | -88.0936
6.3018 | -88.0929
6.3021 | -88.0921
6.3024 | -88.0914
6.3028 | -88.0906
6.3031 | -88.0901
6.3035 | -88.0892
6.3038 | -88.0891
6.3041 | -88.0888
6.3045 | -88.0884
6.3048 | -88.0877
6.3052 | -88.0867
6.3055 | -88.086
6.3058 | -88.0852
6.3062 | -88.0836
6.3065 | -88.0815
6.3069 | -88.0787
6.3072 | -88.0778
6.3075 | -88.0768
6.3079 | -88.0757
6.3082 | -88.0742
6.3086 | -88.0724
6.3089 | -88.0716
6.3092 | -88.0713
6.3096 | -88.0701
6.3099 | -88.0683
6.3103 | -88.0674
6.3106 | -88.066
6.3109 | -88.0638
6.3113 | -88.0626
6.3116 | -88.0625
6.312 | -88.0623
6.3123 | -88.0601
6.3126 | -88.0588
6.313 | -88.0585
6.3133 | -88.0585
6.3137 | -88.0575
6.314 | -88.0579
6.3144 | -88.0579
6.3147 | -88.0569
6.315 | -88.0561
6.3154 | -88.0551
6.3157 | -88.0542
6.3161 | -88.0537
6.3164 | -88.053
6.3167 | -88.0517
6.3171 | -88.0501
6.3174 | -88.0487
6.3178 | -88.0474
6.3181 | -88.0463
6.3184 | -88.0448
6.3188 | -88.0443
6.3191 | -88.0431
6.3195 | -88.0408
6.3198 | -88.0386
6.3201 | -88.0373
6.3205 | -88.0359
6.3208 | -88.0336
6.3212 | -88.031
6.3215 | -88.0284
6.3218 | -88.0266
6.3222 | -88.0243
6.3225 | -88.0204
6.3229 | -88.0181
6.3232 | -88.0167
6.3235 | -88.0154
6.3239 | -88.0137
6.3242 | -88.0115
6.3246 | -88.0089
6.3249 | -88.0067
6.3252 | -88.005
6.3256 | -88.0037
6.3259 | -88.0023
6.3263 | -87.9983
6.3266 | -87.9954
6.327 | -87.9925
6.3273 | -87.9886
6.3276 | -87.9837
6.328 | -87.9805
6.3283 | -87.9779
6.3287 | -87.9758
6.329 | -87.9727
6.3293 | -87.9696
6.3297 | -87.9681
6.33 | -87.9666
6.3304 | -87.9647
6.3307 | -87.9623
6.331 | -87.9592
6.3314 | -87.9566
6.3317 | -87.9544
6.3321 | -87.9517
6.3324 | -87.9492
6.3327 | -87.9472
6.3331 | -87.9452
6.3334 | -87.943
6.3338 | -87.9407
6.3341 | -87.9379
6.3344 | -87.9359
6.3348 | -87.9352
6.3351 | -87.9334
6.3355 | -87.9315
6.3358 | -87.9296
6.3361 | -87.9256
6.3365 | -87.9217
6.3368 | -87.9193
6.3372 | -87.9176
6.3375 | -87.916
6.3378 | -87.9141
6.3382 | -87.9125
6.3385 | -87.9109
6.3389 | -87.909
6.3392 | -87.9072
6.3395 | -87.9054
6.3399 | -87.9031
6.3402 | -87.9009
6.3406 | -87.8994
6.3409 | -87.8977
6.3413 | -87.8957
6.3416 | -87.8935
6.3419 | -87.8914
6.3423 | -87.8899
6.3426 | -87.8881
6.343 | -87.8859
6.3433 | -87.8848
6.3436 | -87.8829
6.344 | -87.8811
6.3443 | -87.8794
6.3447 | -87.8773
6.345 | -87.875
6.3453 | -87.8722
6.3457 | -87.8703
6.346 | -87.8674
6.3464 | -87.8649
6.3467 | -87.8625
6.347 | -87.8608
6.3474 | -87.859
6.3477 | -87.8564
6.3481 | -87.8535
6.3484 | -87.8515
6.3487 | -87.8501
6.3491 | -87.8481
6.3494 | -87.8458
6.3498 | -87.8434
6.3501 | -87.8412
6.3504 | -87.8381
6.3508 | -87.8353
6.3511 | -87.8335
6.3515 | -87.8315
6.3518 | -87.8292
6.3521 | -87.8267
6.3525 | -87.8245
6.3528 | -87.8221
6.3532 | -87.8196
6.3535 | -87.8171
6.3538 | -87.8154
6.3542 | -87.8124
6.3545 | -87.8093
6.3549 | -87.8073
6.3552 | -87.805
6.3556 | -87.8022
6.3559 | -87.7992
6.3562 | -87.7972
6.3566 | -87.7953
6.3569 | -87.7935
6.3573 | -87.7914
6.3576 | -87.7882
6.3579 | -87.7857
6.3583 | -87.7836
6.3586 | -87.7814
6.359 | -87.7795
6.3593 | -87.7774
6.3596 | -87.7752
6.36 | -87.773
6.3603 | -87.7713
6.3607 | -87.7699
6.361 | -87.7681
6.3613 | -87.7655
6.3617 | -87.762
6.362 | -87.7598
6.3624 | -87.7576
6.3627 | -87.755
6.363 | -87.7533
6.3634 | -87.7514
6.3637 | -87.7489
6.3641 | -87.7459
6.3644 | -87.7434
6.3647 | -87.7414
6.3651 | -87.7387
6.3654 | -87.736
6.3658 | -87.7328
6.3661 | -87.7301
6.3664 | -87.7277
6.3668 | -87.7258
6.3671 | -87.7234
6.3675 | -87.7217
6.3678 | -87.7203
6.3681 | -87.7182
6.3685 | -87.7157
6.3688 | -87.7132
6.3692 | -87.7108
6.3695 | -87.7088
6.3699 | -87.7064
6.3702 | -87.7044
6.3705 | -87.7021
6.3709 | -87.6999
6.3712 | -87.6975
6.3716 | -87.6952
6.3719 | -87.6926
6.3722 | -87.6902
6.3726 | -87.6882
6.3729 | -87.6857
6.3733 | -87.6835
6.3736 | -87.6816
6.3739 | -87.6791
6.3743 | -87.6775
6.3746 | -87.6756
6.375 | -87.6735
6.3753 | -87.6718
6.3756 | -87.6668
6.376 | -87.6646
6.3763 | -87.6619
6.3767 | -87.6585
6.377 | -87.6549
6.3773 | -87.6521
6.3777 | -87.6495
6.378 | -87.6469
6.3784 | -87.6442
6.3787 | -87.6413
6.379 | -87.6387
6.3794 | -87.6358
6.3797 | -87.633
6.3801 | -87.6304
6.3804 | -87.6277
6.3807 | -87.6244
6.3811 | -87.6209
6.3814 | -87.6179
6.3818 | -87.6152
6.3821 | -87.6127
6.3825 | -87.6099
6.3828 | -87.6074
6.3831 | -87.605
6.3835 | -87.6025
6.3838 | -87.6001
6.3842 | -87.5975
6.3845 | -87.5955
6.3848 | -87.5931
6.3852 | -87.59
6.3855 | -87.5866
6.3859 | -87.584
6.3862 | -87.5813
6.3865 | -87.5783
6.3869 | -87.5753
6.3872 | -87.5725
6.3876 | -87.5696
6.3879 | -87.567
6.3882 | -87.5645
6.3886 | -87.5618
6.3889 | -87.5591
6.3893 | -87.5565
6.3896 | -87.5542
6.3899 | -87.552
6.3903 | -87.5492
6.3906 | -87.5462
6.391 | -87.5433
6.3913 | -87.5401
6.3916 | -87.5368
6.392 | -87.534
6.3923 | -87.5313
6.3927 | -87.5285
6.393 | -87.5259
6.3933 | -87.5232
6.3937 | -87.5206
6.394 | -87.5177
6.3944 | -87.5152
6.3947 | -87.5128
6.395 | -87.5099
6.3954 | -87.5075
6.3957 | -87.5043
6.3961 | -87.5013
6.3964 | -87.4989
6.3967 | -87.4957
6.3971 | -87.4926
6.3974 | -87.4903
6.3978 | -87.4894
6.3981 | -87.4881
6.3985 | -87.4855
6.3988 | -87.4826
6.3991 | -87.4798
6.3995 | -87.4763
6.3998 | -87.4734
6.4002 | -87.4705
6.4005 | -87.468
6.4008 | -87.4656
6.4012 | -87.463
6.4015 | -87.4603
6.4019 | -87.4574
6.4022 | -87.4544
6.4025 | -87.4527
6.4029 | -87.4504
6.4032 | -87.448
6.4036 | -87.4459
6.4039 | -87.4438
6.4042 | -87.4402
6.4046 | -87.4371
6.4049 | -87.4337
6.4053 | -87.431
6.4056 | -87.4283
6.4059 | -87.4256
6.4063 | -87.423
6.4066 | -87.4209
6.407 | -87.4185
6.4073 | -87.4158
6.4076 | -87.4133
6.408 | -87.4103
6.4083 | -87.4069
6.4087 | -87.4033
6.409 | -87.4005
6.4093 | -87.3981
6.4097 | -87.3956
6.41 | -87.3932
6.4104 | -87.3902
6.4107 | -87.3873
6.4111 | -87.3853
6.4114 | -87.3838
6.4117 | -87.3809
6.4121 | -87.3779
6.4124 | -87.3749
6.4128 | -87.3722
6.4131 | -87.3696
6.4134 | -87.367
6.4138 | -87.3642
6.4141 | -87.3613
6.4145 | -87.3593
6.4148 | -87.3568
6.4151 | -87.356
6.4155 | -87.354
6.4158 | -87.3521
6.4162 | -87.3499
6.4165 | -87.3477
6.4168 | -87.345
6.4172 | -87.3424
6.4175 | -87.3394
6.4179 | -87.3368
6.4182 | -87.334
6.4185 | -87.3318
6.4189 | -87.3295
6.4192 | -87.3271
6.4196 | -87.325
6.4199 | -87.3225
6.4202 | -87.32
6.4206 | -87.3168
6.4209 | -87.3137
6.4213 | -87.3107
6.4216 | -87.3073
6.4219 | -87.3029
6.4223 | -87.2995
6.4226 | -87.2969
6.423 | -87.2939
6.4233 | -87.2918
6.4236 | -87.2897
6.424 | -87.2873
6.4243 | -87.2843
6.4247 | -87.2802
6.425 | -87.2775
6.4254 | -87.2745
6.4257 | -87.2712
6.426 | -87.2685
6.4264 | -87.2658
6.4267 | -87.2631
6.4271 | -87.2606
6.4274 | -87.2576
6.4277 | -87.2544
6.4281 | -87.2513
6.4284 | -87.2485
6.4288 | -87.2457
6.4291 | -87.2432
6.4294 | -87.2411
6.4298 | -87.2389
6.4301 | -87.2363
6.4305 | -87.2336
6.4308 | -87.2302
6.4311 | -87.2264
6.4315 | -87.2239
6.4318 | -87.2207
6.4322 | -87.2176
6.4325 | -87.2149
6.4328 | -87.2124
6.4332 | -87.2098
6.4335 | -87.207
6.4339 | -87.204
6.4342 | -87.2013
6.4345 | -87.1978
6.4349 | -87.1954
6.4352 | -87.193
6.4356 | -87.1909
6.4359 | -87.1881
6.4362 | -87.1861
6.4366 | -87.184
6.4369 | -87.1813
6.4373 | -87.179
6.4376 | -87.1765
6.4379 | -87.1739
6.4383 | -87.1718
6.4386 | -87.1705
6.439 | -87.1704
6.4393 | -87.1686
6.4397 | -87.1663
6.44 | -87.164
6.4403 | -87.1616
6.4407 | -87.1592
6.441 | -87.1566
6.4414 | -87.1539
6.4417 | -87.1507
6.442 | -87.1482
6.4424 | -87.1457
6.4427 | -87.1432
6.4431 | -87.141
6.4434 | -87.1386
6.4437 | -87.1361
6.4441 | -87.1337
6.4444 | -87.1317
6.4448 | -87.129
6.4451 | -87.1271
6.4454 | -87.1251
6.4458 | -87.1225
6.4461 | -87.1203
6.4465 | -87.1182
6.4468 | -87.1156
6.4471 | -87.1127
6.4475 | -87.1099
6.4478 | -87.1079
6.4482 | -87.1054
6.4485 | -87.1035
6.4488 | -87.1012
6.4492 | -87.0988
6.4495 | -87.0962
6.4499 | -87.0934
6.4502 | -87.091
6.4505 | -87.0883
6.4509 | -87.0861
6.4512 | -87.0835
6.4516 | -87.0815
6.4519 | -87.0794
6.4523 | -87.0768
6.4526 | -87.0756
6.4529 | -87.0735
6.4533 | -87.0702
6.4536 | -87.0674
6.454 | -87.0646
6.4543 | -87.0632
6.4546 | -87.061
6.455 | -87.0589
6.4553 | -87.0568
6.4557 | -87.0554
6.456 | -87.0535
6.4563 | -87.0516
6.4567 | -87.0497
6.457 | -87.0479
6.4574 | -87.0456
6.4577 | -87.0434
6.458 | -87.0414
6.4584 | -87.0392
6.4587 | -87.037
6.4591 | -87.0348
6.4594 | -87.0328
6.4597 | -87.0296
6.4601 | -87.0268
6.4604 | -87.0246
6.4608 | -87.0226
6.4611 | -87.0205
6.4614 | -87.0183
6.4618 | -87.0164
6.4621 | -87.0143
6.4625 | -87.0124
6.4628 | -87.0109
6.4631 | -87.009
6.4635 | -87.0069
6.4638 | -87.0057
6.4642 | -87.0036
6.4645 | -87.0015
6.4648 | -86.9993
6.4652 | -86.9972
6.4655 | -86.9953
6.4659 | -86.9933
6.4662 | -86.9913
6.4665 | -86.9899
6.4669 | -86.9879
6.4672 | -86.9848
6.4676 | -86.982
6.4679 | -86.9796
6.4683 | -86.9774
6.4686 | -86.9754
6.4689 | -86.9735
6.4693 | -86.9716
6.4696 | -86.9698
6.47 | -86.9677
6.4703 | -86.9653
6.4706 | -86.9625
6.471 | -86.9608
6.4713 | -86.9592
6.4717 | -86.9575
6.472 | -86.9558
6.4723 | -86.9545
6.4727 | -86.9529
6.473 | -86.9515
6.4734 | -86.9489
6.4737 | -86.9475
6.474 | -86.9454
6.4744 | -86.9437
6.4747 | -86.9419
6.4751 | -86.9401
6.4754 | -86.9382
6.4757 | -86.9372
6.4761 | -86.9356
6.4764 | -86.9342
6.4768 | -86.9329
6.4771 | -86.931
6.4774 | -86.9293
6.4778 | -86.9278
6.4781 | -86.9269
6.4785 | -86.9253
6.4788 | -86.9234
6.4791 | -86.9222
6.4795 | -86.9206
6.4798 | -86.9189
6.4802 | -86.9173
6.4805 | -86.9157
6.4809 | -86.9142
6.4812 | -86.9128
6.4815 | -86.9113
6.4819 | -86.91
6.4822 | -86.9085
6.4826 | -86.9069
6.4829 | -86.9052
6.4832 | -86.9039
6.4836 | -86.9033
6.4839 | -86.902
6.4843 | -86.9005
6.4846 | -86.8995
6.4849 | -86.8989
6.4853 | -86.8987
6.4856 | -86.8979
6.486 | -86.8965
6.4863 | -86.8951
6.4866 | -86.8932
6.487 | -86.8914
6.4873 | -86.8899
6.4877 | -86.8887
6.488 | -86.8873
6.4883 | -86.8861
6.4887 | -86.8851
6.489 | -86.8842
6.4894 | -86.8832
6.4897 | -86.8821
6.49 | -86.8807
6.4904 | -86.8787
6.4907 | -86.8775
6.4911 | -86.8747
6.4914 | -86.8727
6.4917 | -86.8713
6.4921 | -86.8698
6.4924 | -86.8683
6.4928 | -86.8668
6.4931 | -86.8658
6.4934 | -86.8655
6.4938 | -86.8641
6.4941 | -86.8631
6.4945 | -86.862
6.4948 | -86.8607
6.4952 | -86.8594
6.4955 | -86.858
6.4958 | -86.8568
6.4962 | -86.8558
6.4965 | -86.8551
6.4969 | -86.8546
6.4972 | -86.8541
6.4975 | -86.8535
6.4979 | -86.8534
6.4982 | -86.8533
6.4986 | -86.8528
6.4989 | -86.8515
6.4992 | -86.8507
6.4996 | -86.8501
6.4999 | -86.8492
6.5003 | -86.8483
6.5006 | -86.8463
6.5009 | -86.8453
6.5013 | -86.8448
6.5016 | -86.8441
6.502 | -86.8431
6.5023 | -86.8423
6.5026 | -86.8416
6.503 | -86.8407
6.5033 | -86.8415
6.5037 | -86.8409
6.504 | -86.8402
6.5043 | -86.8392
6.5047 | -86.8385
6.505 | -86.8384
6.5054 | -86.8387
6.5057 | -86.8392
6.506 | -86.8387
6.5064 | -86.8378
6.5067 | -86.837
6.5071 | -86.8364
6.5074 | -86.8356
6.5077 | -86.835
6.5081 | -86.8344
6.5084 | -86.8335
6.5088 | -86.833
6.5091 | -86.8325
6.5095 | -86.8318
6.5098 | -86.8306
6.5101 | -86.8303
6.5105 | -86.8298
6.5108 | -86.8287
6.5112 | -86.828
6.5115 | -86.8269
6.5118 | -86.8251
6.5122 | -86.8244
6.5125 | -86.8241
6.5129 | -86.8236
6.5132 | -86.8248
6.5135 | -86.8239
6.5139 | -86.8235
6.5142 | -86.8226
6.5146 | -86.8222
6.5149 | -86.8221
6.5152 | -86.8219
6.5156 | -86.8217
6.5159 | -86.8215
6.5163 | -86.8211
6.5166 | -86.8208
6.5169 | -86.8192
6.5173 | -86.8172
6.5176 | -86.8145
6.518 | -86.8132
6.5183 | -86.8121
6.5186 | -86.8112
6.519 | -86.8109
6.5193 | -86.8108
6.5197 | -86.8102
6.52 | -86.8095
6.5203 | -86.8092
6.5207 | -86.8092
6.521 | -86.8098
6.5214 | -86.8095
6.5217 | -86.8091
6.5221 | -86.809
6.5224 | -86.8089
6.5227 | -86.8084
6.5231 | -86.8082
6.5234 | -86.8086
6.5238 | -86.8087
6.5241 | -86.8087
6.5244 | -86.8088
6.5248 | -86.8084
6.5251 | -86.8083
6.5255 | -86.8088
6.5258 | -86.8086
6.5261 | -86.8083
6.5265 | -86.8083
6.5268 | -86.8083
6.5272 | -86.808
6.5275 | -86.8077
6.5278 | -86.8078
6.5282 | -86.808
6.5285 | -86.808
6.5289 | -86.808
6.5292 | -86.8078
6.5295 | -86.8078
6.5299 | -86.808
6.5302 | -86.8089
6.5306 | -86.8092
6.5309 | -86.8095
6.5312 | -86.81
6.5316 | -86.8102
6.5319 | -86.8099
6.5323 | -86.8093
6.5326 | -86.811
6.5329 | -86.8111
6.5333 | -86.8114
6.5336 | -86.8119
6.534 | -86.8117
6.5343 | -86.8112
6.5346 | -86.8115
6.535 | -86.8111
6.5353 | -86.8106
6.5357 | -86.8111
6.536 | -86.8107
6.5364 | -86.811
6.5367 | -86.8112
6.537 | -86.8114
6.5374 | -86.8116
6.5377 | -86.8118
6.5381 | -86.8114
6.5384 | -86.8107
6.5387 | -86.8105
6.5391 | -86.8106
6.5394 | -86.8116
6.5398 | -86.8124
6.5401 | -86.8128
6.5404 | -86.8132
6.5408 | -86.8137
6.5411 | -86.8147
6.5415 | -86.8153
6.5418 | -86.8154
6.5421 | -86.8157
6.5425 | -86.8179
6.5428 | -86.8182
6.5432 | -86.8189
6.5435 | -86.82
6.5438 | -86.8207
6.5442 | -86.8212
6.5445 | -86.8217
6.5449 | -86.8222
6.5452 | -86.8229
6.5455 | -86.8234
6.5459 | -86.8236
6.5462 | -86.8239
6.5466 | -86.8244
6.5469 | -86.8247
6.5472 | -86.8253
6.5476 | -86.8265
6.5479 | -86.8273
6.5483 | -86.8271
6.5486 | -86.8277
6.5489 | -86.8283
6.5493 | -86.8289
6.5496 | -86.8292
6.55 | -86.8289
6.5503 | -86.8295
6.5507 | -86.8298
6.551 | -86.8296
6.5513 | -86.8309
6.5517 | -86.832
6.552 | -86.8329
6.5524 | -86.8352
6.5527 | -86.8362
6.553 | -86.8371
6.5534 | -86.8377
6.5537 | -86.8385
6.5541 | -86.8388
6.5544 | -86.8395
6.5547 | -86.84
6.5551 | -86.8407
6.5554 | -86.8416
6.5558 | -86.8424
6.5561 | -86.8431
6.5564 | -86.8439
6.5568 | -86.845
6.5571 | -86.846
6.5575 | -86.8469
6.5578 | -86.8471
6.5581 | -86.8478
6.5585 | -86.8488
6.5588 | -86.8504
6.5592 | -86.8522
6.5595 | -86.8531
6.5598 | -86.8538
6.5602 | -86.8549
6.5605 | -86.8561
6.5609 | -86.8572
6.5612 | -86.8582
6.5615 | -86.8592
6.5619 | -86.8609
6.5622 | -86.8634
6.5626 | -86.8646
6.5629 | -86.8653
6.5632 | -86.8661
6.5636 | -86.8668
6.5639 | -86.8678
6.5643 | -86.8688
6.5646 | -86.8699
6.565 | -86.8711
6.5653 | -86.8722
6.5656 | -86.8727
6.566 | -86.8735
6.5663 | -86.8743
6.5667 | -86.8753
6.567 | -86.8767
6.5673 | -86.8773
6.5677 | -86.8778
6.568 | -86.8791
6.5684 | -86.8805
6.5687 | -86.8818
6.569 | -86.8832
6.5694 | -86.8843
6.5697 | -86.885
6.5701 | -86.8864
6.5704 | -86.8883
6.5707 | -86.8897
6.5711 | -86.8911
6.5714 | -86.8926
6.5718 | -86.8939
6.5721 | -86.8962
6.5724 | -86.8976
6.5728 | -86.8991
6.5731 | -86.9013
6.5735 | -86.9035
6.5738 | -86.9051
6.5741 | -86.9061
6.5745 | -86.9075
6.5748 | -86.9088
6.5752 | -86.9107
6.5755 | -86.9126
6.5758 | -86.9138
6.5762 | -86.9149
6.5765 | -86.9163
6.5769 | -86.9179
6.5772 | -86.919
6.5775 | -86.9204
6.5779 | -86.9217
6.5782 | -86.9231
6.5786 | -86.9243
6.5789 | -86.926
6.5793 | -86.928
6.5796 | -86.9296
6.5799 | -86.931
6.5803 | -86.9326
6.5806 | -86.9341
6.581 | -86.9357
6.5813 | -86.9375
6.5816 | -86.9391
6.582 | -86.9419
6.5823 | -86.9431
6.5827 | -86.9453
6.583 | -86.9474
6.5833 | -86.9489
6.5837 | -86.9504
6.584 | -86.9522
6.5844 | -86.9542
6.5847 | -86.9559
6.585 | -86.9567
6.5854 | -86.9579
6.5857 | -86.9599
6.5861 | -86.9614
6.5864 | -86.9629
6.5867 | -86.9642
6.5871 | -86.9661
6.5874 | -86.9684
6.5878 | -86.9706
6.5881 | -86.9727
6.5884 | -86.9746
6.5888 | -86.9766
6.5891 | -86.9784
6.5895 | -86.9803
6.5898 | -86.9817
6.5901 | -86.9836
6.5905 | -86.9855
6.5908 | -86.9876
6.5912 | -86.9888
6.5915 | -86.9914
6.5918 | -86.9931
6.5922 | -86.9949
6.5925 | -86.9973
6.5929 | -86.9993
6.5932 | -87.0003
6.5936 | -87.0023
6.5939 | -87.0048
6.5942 | -87.0065
6.5946 | -87.0083
6.5949 | -87.0097
6.5953 | -87.0107
6.5956 | -87.0127
6.5959 | -87.0146
6.5963 | -87.0174
6.5966 | -87.0191
6.597 | -87.0208
6.5973 | -87.0224
6.5976 | -87.0241
6.598 | -87.0265
6.5983 | -87.0291
6.5987 | -87.0315
6.599 | -87.0335
6.5993 | -87.0352
6.5997 | -87.037
6.6 | -87.0389
6.6004 | -87.0405
6.6007 | -87.0419
6.601 | -87.0436
6.6014 | -87.0461
6.6017 | -87.0481
6.6021 | -87.05
6.6024 | -87.052
6.6027 | -87.054
6.6031 | -87.0557
6.6034 | -87.0575
6.6038 | -87.0591
6.6041 | -87.0622
6.6044 | -87.066
6.6048 | -87.0688
6.6051 | -87.0709
6.6055 | -87.0731
6.6058 | -87.0756
6.6062 | -87.0772
6.6065 | -87.0801
6.6068 | -87.0844
6.6072 | -87.0874
6.6075 | -87.0901
6.6079 | -87.0924
6.6082 | -87.0942
6.6085 | -87.0963
6.6089 | -87.0977
6.6092 | -87.0987
6.6096 | -87.1004
6.6099 | -87.1026
6.6102 | -87.1045
6.6106 | -87.1051
6.6109 | -87.1069
6.6113 | -87.1089
6.6116 | -87.1109
6.6119 | -87.113
6.6123 | -87.115
6.6126 | -87.117
6.613 | -87.1188
6.6133 | -87.1208
6.6136 | -87.1229
6.614 | -87.1251
6.6143 | -87.1271
6.6147 | -87.1288
6.615 | -87.1306
6.6153 | -87.1331
6.6157 | -87.1355
6.616 | -87.1376
6.6164 | -87.1397
6.6167 | -87.1417
6.617 | -87.1438
6.6174 | -87.1459
6.6177 | -87.148
6.6181 | -87.1502
6.6184 | -87.1521
6.6187 | -87.1535
6.6191 | -87.1542
6.6194 | -87.1557
6.6198 | -87.1579
6.6201 | -87.1598
6.6205 | -87.1616
6.6208 | -87.1629
6.6211 | -87.1654
6.6215 | -87.1678
6.6218 | -87.17
6.6222 | -87.1721
6.6225 | -87.1737
6.6228 | -87.1755
6.6232 | -87.1776
6.6235 | -87.1797
6.6239 | -87.1815
6.6242 | -87.1842
6.6245 | -87.1871
6.6249 | -87.189
6.6252 | -87.1909
6.6256 | -87.1929
6.6259 | -87.1956
6.6262 | -87.1982
6.6266 | -87.2004
6.6269 | -87.2024
6.6273 | -87.2044
6.6276 | -87.2062
6.6279 | -87.2083
6.6283 | -87.21
6.6286 | -87.2121
6.629 | -87.2142
6.6293 | -87.2164
6.6296 | -87.2187
6.63 | -87.2209
6.6303 | -87.223
6.6307 | -87.2251
6.631 | -87.2276
6.6313 | -87.2296
6.6317 | -87.2317
6.632 | -87.2335
6.6324 | -87.2352
6.6327 | -87.2365
6.633 | -87.2383
6.6334 | -87.2404
6.6337 | -87.2426
6.6341 | -87.2448
6.6344 | -87.247
6.6348 | -87.2492
6.6351 | -87.2511
6.6354 | -87.2534
6.6358 | -87.256
6.6361 | -87.2579
6.6365 | -87.2601
6.6368 | -87.2625
6.6371 | -87.266
6.6375 | -87.2691
6.6378 | -87.2711
6.6382 | -87.273
6.6385 | -87.2748
6.6388 | -87.2773
6.6392 | -87.2815
6.6395 | -87.2843
6.6399 | -87.2862
6.6402 | -87.2877
6.6405 | -87.2894
6.6409 | -87.2919
6.6412 | -87.2937
6.6416 | -87.2956
6.6419 | -87.2976
6.6422 | -87.2996
6.6426 | -87.3015
6.6429 | -87.3037
6.6433 | -87.3054
6.6436 | -87.308
6.6439 | -87.3108
6.6443 | -87.3128
6.6446 | -87.3148
6.645 | -87.3169
6.6453 | -87.3191
6.6456 | -87.3212
6.646 | -87.3236
6.6463 | -87.3258
6.6467 | -87.3278
6.647 | -87.3303
6.6474 | -87.3321
6.6477 | -87.3341
6.648 | -87.3365
6.6484 | -87.3387
6.6487 | -87.3409
6.6491 | -87.3434
6.6494 | -87.3462
6.6497 | -87.3484
6.6501 | -87.3507
6.6504 | -87.3525
6.6508 | -87.3548
6.6511 | -87.3568
6.6514 | -87.359
6.6518 | -87.3613
6.6521 | -87.3636
6.6525 | -87.3654
6.6528 | -87.3678
6.6531 | -87.37
6.6535 | -87.3724
6.6538 | -87.3748
6.6542 | -87.3769
6.6545 | -87.3791
6.6548 | -87.3814
6.6552 | -87.3837
6.6555 | -87.3864
6.6559 | -87.3889
6.6562 | -87.3906
6.6565 | -87.3917
6.6569 | -87.3937
6.6572 | -87.3957
6.6576 | -87.398
6.6579 | -87.3993
6.6582 | -87.4005
6.6586 | -87.4024
6.6589 | -87.4043
6.6593 | -87.4063
6.6596 | -87.4084
6.6599 | -87.4106
6.6603 | -87.4126
6.6606 | -87.4142
6.661 | -87.4164
6.6613 | -87.4191
6.6616 | -87.4214
6.662 | -87.4235
6.6623 | -87.4258
6.6627 | -87.4281
6.663 | -87.4303
6.6634 | -87.4328
6.6637 | -87.4347
6.664 | -87.4364
6.6644 | -87.4386
6.6647 | -87.4407
6.6651 | -87.4429
6.6654 | -87.4451
6.6657 | -87.4469
6.6661 | -87.4484
6.6664 | -87.4503
6.6668 | -87.4525
6.6671 | -87.4548
6.6674 | -87.4566
6.6678 | -87.4589
6.6681 | -87.4606
6.6685 | -87.4627
6.6688 | -87.4646
6.6691 | -87.4668
6.6695 | -87.4691
6.6698 | -87.4711
6.6702 | -87.4733
6.6705 | -87.4757
6.6708 | -87.4779
6.6712 | -87.4809
6.6715 | -87.484
6.6719 | -87.4865
6.6722 | -87.4888
6.6725 | -87.4909
6.6729 | -87.4928
6.6732 | -87.4949
6.6736 | -87.4971
6.6739 | -87.4992
6.6742 | -87.502
6.6746 | -87.5049
6.6749 | -87.5072
6.6753 | -87.5093
6.6756 | -87.5113
6.676 | -87.5133
6.6763 | -87.5152
6.6766 | -87.5175
6.677 | -87.5202
6.6773 | -87.5223
6.6777 | -87.5242
6.678 | -87.5263
6.6783 | -87.529
6.6787 | -87.5316
6.679 | -87.534
6.6794 | -87.5362
6.6797 | -87.5384
6.68 | -87.5403
6.6804 | -87.5406
6.6807 | -87.5412
6.6811 | -87.5433
6.6814 | -87.5451
6.6817 | -87.5468
6.6821 | -87.5498
6.6824 | -87.5533
6.6828 | -87.5565
6.6831 | -87.5592
6.6834 | -87.5623
6.6838 | -87.5645
6.6841 | -87.5666
6.6845 | -87.5682
6.6848 | -87.5706
6.6851 | -87.5736
6.6855 | -87.5758
6.6858 | -87.5777
6.6862 | -87.5799
6.6865 | -87.5819
6.6868 | -87.584
6.6872 | -87.5862
6.6875 | -87.5882
6.6879 | -87.5903
6.6882 | -87.5921
6.6885 | -87.5944
6.6889 | -87.5966
6.6892 | -87.5992
6.6896 | -87.6034
6.6899 | -87.6056
6.6903 | -87.6088
6.6906 | -87.6108
6.6909 | -87.613
6.6913 | -87.615
6.6916 | -87.6163
6.692 | -87.6181
6.6923 | -87.6199
6.6926 | -87.6223
6.693 | -87.6241
6.6933 | -87.6259
6.6937 | -87.6283
6.694 | -87.6302
6.6943 | -87.6326
6.6947 | -87.635
6.695 | -87.6377
6.6954 | -87.6396
6.6957 | -87.6416
6.696 | -87.6436
6.6964 | -87.6456
6.6967 | -87.6478
6.6971 | -87.6506
6.6974 | -87.6536
6.6977 | -87.6558
6.6981 | -87.6574
6.6984 | -87.6592
6.6988 | -87.6623
6.6991 | -87.6658
6.6994 | -87.6684
6.6998 | -87.6704
6.7001 | -87.6729
6.7005 | -87.675
6.7008 | -87.6765
6.7011 | -87.6786
6.7015 | -87.6816
6.7018 | -87.6839
6.7022 | -87.6858
6.7025 | -87.6879
6.7028 | -87.6904
6.7032 | -87.6925
6.7035 | -87.6939
6.7039 | -87.6957
6.7042 | -87.6968
6.7046 | -87.6986
6.7049 | -87.7015
6.7052 | -87.7042
6.7056 | -87.7063
6.7059 | -87.7092
6.7063 | -87.7111
6.7066 | -87.7127
6.7069 | -87.7152
6.7073 | -87.7176
6.7076 | -87.7203
6.708 | -87.7227
6.7083 | -87.725
6.7086 | -87.7266
6.709 | -87.7282
6.7093 | -87.7302
6.7097 | -87.7321
6.71 | -87.734
6.7103 | -87.7355
6.7107 | -87.7367
6.711 | -87.7385
6.7114 | -87.7406
6.7117 | -87.7422
6.712 | -87.7436
6.7124 | -87.7446
6.7127 | -87.7465
6.7131 | -87.7485
6.7134 | -87.7509
6.7137 | -87.7529
6.7141 | -87.7552
6.7144 | -87.7571
6.7148 | -87.7589
6.7151 | -87.761
6.7154 | -87.7628
6.7158 | -87.7645
6.7161 | -87.7667
6.7165 | -87.7693
6.7168 | -87.7709
6.7172 | -87.7728
6.7175 | -87.7748
6.7178 | -87.7768
6.7182 | -87.7779
6.7185 | -87.7808
6.7189 | -87.7829
6.7192 | -87.7849
6.7195 | -87.7857
6.7199 | -87.7877
6.7202 | -87.7901
6.7206 | -87.7923
6.7209 | -87.7955
6.7212 | -87.799
6.7216 | -87.8026
6.7219 | -87.8065
6.7223 | -87.8105
6.7226 | -87.8135
6.7229 | -87.8165
6.7233 | -87.8194
6.7236 | -87.8217
6.724 | -87.824
6.7243 | -87.8272
6.7246 | -87.8313
6.725 | -87.8329
6.7253 | -87.8339
6.7257 | -87.8358
6.726 | -87.8376
6.7263 | -87.8394
6.7267 | -87.8413
6.727 | -87.8432
6.7274 | -87.8454
6.7277 | -87.8473
6.728 | -87.849
6.7284 | -87.8509
6.7287 | -87.8529
6.7291 | -87.8545
6.7294 | -87.856
6.7297 | -87.8576
6.7301 | -87.8585
6.7304 | -87.859
6.7308 | -87.8605
6.7311 | -87.8618
6.7314 | -87.8641
6.7318 | -87.8659
6.7321 | -87.8667
6.7325 | -87.8683
6.7328 | -87.8705
6.7332 | -87.8719
6.7335 | -87.8734
6.7338 | -87.8749
6.7342 | -87.878
6.7345 | -87.881
6.7349 | -87.8824
6.7352 | -87.8838
6.7355 | -87.8852
6.7359 | -87.8868
6.7362 | -87.8882
6.7366 | -87.8899
6.7369 | -87.8914
6.7372 | -87.8932
6.7376 | -87.8944
6.7379 | -87.896
6.7383 | -87.8985
6.7386 | -87.9002
6.7389 | -87.9011
6.7393 | -87.9029
6.7396 | -87.9046
6.74 | -87.9062
6.7403 | -87.908
6.7406 | -87.9098
6.741 | -87.9115
6.7413 | -87.9133
6.7417 | -87.9151
6.742 | -87.9173
6.7423 | -87.9203
6.7427 | -87.9222
6.743 | -87.9241
6.7434 | -87.9263
6.7437 | -87.9279
6.744 | -87.9295
6.7444 | -87.9311
6.7447 | -87.9325
6.7451 | -87.9351
6.7454 | -87.9384
6.7458 | -87.9413
6.7461 | -87.9436
6.7464 | -87.9457
6.7468 | -87.9476
6.7471 | -87.9493
6.7475 | -87.9515
6.7478 | -87.9533
6.7481 | -87.9553
6.7485 | -87.9574
6.7488 | -87.9572
6.7492 | -87.9588
6.7495 | -87.9601
6.7498 | -87.9613
6.7502 | -87.963
6.7505 | -87.9646
6.7509 | -87.9665
6.7512 | -87.9686
6.7515 | -87.9707
6.7519 | -87.972
6.7522 | -87.9731
6.7526 | -87.975
6.7529 | -87.9767
6.7532 | -87.9781
6.7536 | -87.9801
6.7539 | -87.9821
6.7543 | -87.9839
6.7546 | -87.9853
6.7549 | -87.9868
6.7553 | -87.9884
6.7556 | -87.99
6.756 | -87.9915
6.7563 | -87.9932
6.7566 | -87.9949
6.757 | -87.9968
6.7573 | -87.9984
6.7577 | -87.9998
6.758 | -88.001
6.7583 | -88.0027
6.7587 | -88.0043
6.759 | -88.0056
6.7594 | -88.0072
6.7597 | -88.0093
6.7601 | -88.0118
6.7604 | -88.0143
6.7607 | -88.016
6.7611 | -88.0174
6.7614 | -88.0185
6.7618 | -88.0197
6.7621 | -88.0214
6.7624 | -88.0238
6.7628 | -88.0257
6.7631 | -88.0274
6.7635 | -88.03
6.7638 | -88.0329
6.7641 | -88.0343
6.7645 | -88.0359
6.7648 | -88.0384
6.7652 | -88.0399
6.7655 | -88.0413
6.7658 | -88.0432
6.7662 | -88.0459
6.7665 | -88.0479
6.7669 | -88.0493
6.7672 | -88.0506
6.7675 | -88.052
6.7679 | -88.0534
6.7682 | -88.0548
6.7686 | -88.0562
6.7689 | -88.0574
6.7692 | -88.0586
6.7696 | -88.06
6.7699 | -88.0615
6.7703 | -88.0629
6.7706 | -88.0645
6.7709 | -88.0661
6.7713 | -88.0678
6.7716 | -88.0699
6.772 | -88.0718
6.7723 | -88.0738
6.7726 | -88.0756
6.773 | -88.0773
6.7733 | -88.0787
6.7737 | -88.0807
6.774 | -88.0836
6.7744 | -88.0863
6.7747 | -88.0876
6.775 | -88.089
6.7754 | -88.0917
6.7757 | -88.0947
6.7761 | -88.0975
6.7764 | -88.0996
6.7767 | -88.1021
6.7771 | -88.1058
6.7774 | -88.1087
6.7778 | -88.1102
6.7781 | -88.1116
6.7784 | -88.1135
6.7788 | -88.1153
6.7791 | -88.1178
6.7795 | -88.1207
6.7798 | -88.1236
6.7801 | -88.1255
6.7805 | -88.1262
6.7808 | -88.1267
6.7812 | -88.1289
6.7815 | -88.132
6.7818 | -88.135
6.7822 | -88.1378
6.7825 | -88.1409
6.7829 | -88.1449
6.7832 | -88.1484
6.7835 | -88.1509
6.7839 | -88.1533
6.7842 | -88.1564
6.7846 | -88.1591
6.7849 | -88.1607
6.7852 | -88.1624
6.7856 | -88.1638
6.7859 | -88.1656
6.7863 | -88.1681
6.7866 | -88.1706
6.787 | -88.1725
6.7873 | -88.1749
6.7876 | -88.1771
6.788 | -88.1785
6.7883 | -88.1802
6.7887 | -88.1818
6.789 | -88.1839
6.7893 | -88.1863
6.7897 | -88.1885
6.79 | -88.1902
6.7904 | -88.1915
6.7907 | -88.1931
6.791 | -88.1949
6.7914 | -88.1965
6.7917 | -88.1982
6.7921 | -88.2005
6.7924 | -88.2033
6.7927 | -88.2054
6.7931 | -88.2065
6.7934 | -88.2079
6.7938 | -88.2094
6.7941 | -88.2108
6.7944 | -88.2123
6.7948 | -88.2136
6.7951 | -88.2147
6.7955 | -88.2161
6.7958 | -88.2179
6.7961 | -88.2196
6.7965 | -88.2208
6.7968 | -88.2216
6.7972 | -88.2227
6.7975 | -88.2244
6.7978 | -88.2263
6.7982 | -88.2278
6.7985 | -88.2293
6.7989 | -88.2313
6.7992 | -88.233
6.7995 | -88.234
6.7999 | -88.235
6.8002 | -88.2364
6.8006 | -88.2378
6.8009 | -88.2393
6.8012 | -88.241
6.8016 | -88.2433
6.8019 | -88.2448
6.8023 | -88.246
6.8026 | -88.247
6.803 | -88.2488
6.8033 | -88.2506
6.8036 | -88.252
6.804 | -88.2534
6.8043 | -88.2554
6.8047 | -88.2576
6.805 | -88.2593
6.8053 | -88.2607
6.8057 | -88.2621
6.806 | -88.2637
6.8064 | -88.2657
6.8067 | -88.2674
6.807 | -88.2699
6.8074 | -88.2724
6.8077 | -88.2744
6.8081 | -88.2759
6.8084 | -88.2773
6.8087 | -88.279
6.8091 | -88.2814
6.8094 | -88.283
6.8098 | -88.2846
6.8101 | -88.2861
6.8104 | -88.288
6.8108 | -88.2896
6.8111 | -88.2922
6.8115 | -88.2946
6.8118 | -88.2957
6.8121 | -88.2977
6.8125 | -88.2997
6.8128 | -88.3015
6.8132 | -88.303
6.8135 | -88.3039
6.8138 | -88.3044
6.8142 | -88.3052
6.8145 | -88.3071
6.8149 | -88.3093
6.8152 | -88.3107
6.8156 | -88.3122
6.8159 | -88.3132
6.8162 | -88.3136
6.8166 | -88.3149
6.8169 | -88.3163
6.8173 | -88.3178
6.8176 | -88.3196
6.8179 | -88.3204
6.8183 | -88.3205
6.8186 | -88.3215
6.819 | -88.3228
6.8193 | -88.3254
6.8196 | -88.3269
6.82 | -88.3283
6.8203 | -88.33
6.8207 | -88.332
6.821 | -88.3337
6.8213 | -88.3356
6.8217 | -88.3379
6.822 | -88.3397
6.8224 | -88.3412
6.8227 | -88.3428
6.823 | -88.3448
6.8234 | -88.3464
6.8237 | -88.3483
6.8241 | -88.3499
6.8244 | -88.351
6.8247 | -88.3521
6.8251 | -88.3535
6.8254 | -88.3549
6.8258 | -88.3563
6.8261 | -88.3578
6.8264 | -88.3594
6.8268 | -88.3611
6.8271 | -88.3626
6.8275 | -88.3643
6.8278 | -88.3654
6.8281 | -88.3677
6.8285 | -88.3693
6.8288 | -88.3706
6.8292 | -88.3718
6.8295 | -88.373
6.8299 | -88.3742
6.8302 | -88.3757
6.8305 | -88.3772
6.8309 | -88.3788
6.8312 | -88.3798
6.8316 | -88.3812
6.8319 | -88.3825
6.8322 | -88.3846
6.8326 | -88.3862
6.8329 | -88.3878
6.8333 | -88.3889
6.8336 | -88.3902
6.8339 | -88.3932
6.8343 | -88.3961
6.8346 | -88.3977
6.835 | -88.3989
6.8353 | -88.4014
6.8356 | -88.4043
6.836 | -88.4064
6.8363 | -88.408
6.8367 | -88.4096
6.837 | -88.4113
6.8373 | -88.4134
6.8377 | -88.4151
6.838 | -88.416
6.8384 | -88.4182
6.8387 | -88.4201
6.839 | -88.4217
6.8394 | -88.4237
6.8397 | -88.426
6.8401 | -88.4285
6.8404 | -88.4304
6.8407 | -88.4321
6.8411 | -88.4328
6.8414 | -88.4337
6.8418 | -88.4342
6.8421 | -88.4352
6.8424 | -88.4366
6.8428 | -88.438
6.8431 | -88.4396
6.8435 | -88.4411
6.8438 | -88.4429
6.8442 | -88.445
6.8445 | -88.4471
6.8448 | -88.4484
6.8452 | -88.4498
6.8455 | -88.4514
6.8459 | -88.4533
6.8462 | -88.4553
6.8465 | -88.4571
6.8469 | -88.4586
6.8472 | -88.4609
6.8476 | -88.4625
6.8479 | -88.4641
6.8482 | -88.4656
6.8486 | -88.4676
6.8489 | -88.4693
6.8493 | -88.4708
6.8496 | -88.4725
6.8499 | -88.4744
6.8503 | -88.4765
6.8506 | -88.4784
6.851 | -88.4808
6.8513 | -88.483
6.8516 | -88.4836
6.852 | -88.485
6.8523 | -88.4877
6.8527 | -88.4897
6.853 | -88.4912
6.8533 | -88.493
6.8537 | -88.4949
6.854 | -88.4963
6.8544 | -88.4974
6.8547 | -88.4986
6.855 | -88.5
6.8554 | -88.5019
6.8557 | -88.5049
6.8561 | -88.507
6.8564 | -88.5087
6.8567 | -88.5105
6.8571 | -88.5122
6.8574 | -88.5144
6.8578 | -88.5153
6.8581 | -88.5163
6.8585 | -88.5178
6.8588 | -88.5196
6.8591 | -88.5212
6.8595 | -88.5223
6.8598 | -88.5249
6.8602 | -88.5267
6.8605 | -88.5289
6.8608 | -88.5307
6.8612 | -88.5325
6.8615 | -88.5342
6.8619 | -88.5357
6.8622 | -88.5375
6.8625 | -88.5399
6.8629 | -88.5417
6.8632 | -88.544
6.8636 | -88.5458
6.8639 | -88.5472
6.8642 | -88.5486
6.8646 | -88.5516
6.8649 | -88.5539
6.8653 | -88.556
6.8656 | -88.5581
6.8659 | -88.559
6.8663 | -88.5604
6.8666 | -88.5613
6.867 | -88.5631
6.8673 | -88.5648
6.8676 | -88.5663
6.868 | -88.568
6.8683 | -88.5697
6.8687 | -88.5716
6.869 | -88.5736
6.8693 | -88.576
6.8697 | -88.5783
6.87 | -88.5799
6.8704 | -88.5818
6.8707 | -88.5839
6.8711 | -88.5858
6.8714 | -88.5879
6.8717 | -88.5897
6.8721 | -88.5913
6.8724 | -88.5929
6.8728 | -88.5949
6.8731 | -88.5971
6.8734 | -88.5988
6.8738 | -88.6007
6.8741 | -88.6024
6.8745 | -88.6038
6.8748 | -88.6053
6.8751 | -88.6075
6.8755 | -88.6094
6.8758 | -88.611
6.8762 | -88.6132
6.8765 | -88.6157
6.8768 | -88.6178
6.8772 | -88.6197
6.8775 | -88.6209
6.8779 | -88.623
6.8782 | -88.6249
6.8785 | -88.6267
6.8789 | -88.6287
6.8792 | -88.6311
6.8796 | -88.6327
6.8799 | -88.6344
6.8802 | -88.6362
6.8806 | -88.6378
6.8809 | -88.6399
6.8813 | -88.6416
6.8816 | -88.6435
6.8819 | -88.6455
6.8823 | -88.648
6.8826 | -88.6504
6.883 | -88.6528
6.8833 | -88.6554
6.8836 | -88.6593
6.884 | -88.6627
6.8843 | -88.6652
6.8847 | -88.6676
6.885 | -88.6699
6.8854 | -88.6718
6.8857 | -88.6737
6.886 | -88.6764
6.8864 | -88.679
6.8867 | -88.6825
6.8871 | -88.6854
6.8874 | -88.6872
6.8877 | -88.6894
6.8881 | -88.6917
6.8884 | -88.6936
6.8888 | -88.6955
6.8891 | -88.6974
6.8894 | -88.6991
6.8898 | -88.7017
6.8901 | -88.7046
6.8905 | -88.7071
6.8908 | -88.7094
6.8911 | -88.7112
6.8915 | -88.7132
6.8918 | -88.7152
6.8922 | -88.7173
6.8925 | -88.7195
6.8928 | -88.7213
6.8932 | -88.724
6.8935 | -88.728
6.8939 | -88.7311
6.8942 | -88.7341
6.8945 | -88.7379
6.8949 | -88.7406
6.8952 | -88.7423
6.8956 | -88.7439
6.8959 | -88.746
6.8962 | -88.7485
6.8966 | -88.7508
6.8969 | -88.7527
6.8973 | -88.7551
6.8976 | -88.7576
6.8979 | -88.7595
6.8983 | -88.7608
6.8986 | -88.7623
6.899 | -88.7643
6.8993 | -88.7659
6.8997 | -88.7676
6.9 | -88.7691
6.9003 | -88.7707
6.9007 | -88.7724
6.901 | -88.7744
6.9014 | -88.7763
6.9017 | -88.7786
6.902 | -88.7805
6.9024 | -88.7824
6.9027 | -88.7842
6.9031 | -88.7861
6.9034 | -88.7881
6.9037 | -88.7901
6.9041 | -88.7921
6.9044 | -88.7943
6.9048 | -88.7969
6.9051 | -88.798
6.9054 | -88.7992
6.9058 | -88.801
6.9061 | -88.8029
6.9065 | -88.8045
6.9068 | -88.8064
6.9071 | -88.8085
6.9075 | -88.8103
6.9078 | -88.8115
6.9082 | -88.8134
6.9085 | -88.8158
6.9088 | -88.818
6.9092 | -88.8201
6.9095 | -88.8218
6.9099 | -88.8232
6.9102 | -88.8241
6.9105 | -88.8257
6.9109 | -88.8281
6.9112 | -88.8302
6.9116 | -88.8319
6.9119 | -88.8341
6.9122 | -88.8361
6.9126 | -88.8382
6.9129 | -88.8401
6.9133 | -88.8422
6.9136 | -88.8446
6.914 | -88.8464
6.9143 | -88.8486
6.9146 | -88.8508
6.915 | -88.8526
6.9153 | -88.855
6.9157 | -88.8569
6.916 | -88.8595
6.9163 | -88.862
6.9167 | -88.864
6.917 | -88.8661
6.9174 | -88.8694
6.9177 | -88.8724
6.918 | -88.8744
6.9184 | -88.8763
6.9187 | -88.8784
6.9191 | -88.8804
6.9194 | -88.8821
6.9197 | -88.8841
6.9201 | -88.8858
6.9204 | -88.8888
6.9208 | -88.8908
6.9211 | -88.8929
6.9214 | -88.8948
6.9218 | -88.8967
6.9221 | -88.8986
6.9225 | -88.9006
6.9228 | -88.9027
6.9231 | -88.9047
6.9235 | -88.9068
6.9238 | -88.9087
6.9242 | -88.9123
6.9245 | -88.9156
6.9248 | -88.9178
6.9252 | -88.9196
6.9255 | -88.922
6.9259 | -88.924
6.9262 | -88.9262
6.9265 | -88.9281
6.9269 | -88.9304
6.9272 | -88.9325
6.9276 | -88.9346
6.9279 | -88.9367
6.9283 | -88.9388
6.9286 | -88.9413
6.9289 | -88.9438
6.9293 | -88.9458
6.9296 | -88.9478
6.93 | -88.9504
6.9303 | -88.9525
6.9306 | -88.9539
6.931 | -88.9563
6.9313 | -88.9593
6.9317 | -88.9616
6.932 | -88.9637
6.9323 | -88.9675
6.9327 | -88.971
6.933 | -88.9739
6.9334 | -88.9768
6.9337 | -88.979
6.934 | -88.9808
6.9344 | -88.9824
6.9347 | -88.9844
6.9351 | -88.9869
6.9354 | -88.9884
6.9357 | -88.9914
6.9361 | -88.9949
6.9364 | -88.9971
6.9368 | -88.999
6.9371 | -89.0013
6.9374 | -89.0034
6.9378 | -89.0056
6.9381 | -89.0078
6.9385 | -89.01
6.9388 | -89.0137
6.9391 | -89.0171
6.9395 | -89.0193
6.9398 | -89.0216
6.9402 | -89.0235
6.9405 | -89.0256
6.9409 | -89.0281
6.9412 | -89.0311
6.9415 | -89.0332
6.9419 | -89.0347
6.9422 | -89.0365
6.9426 | -89.0388
6.9429 | -89.0417
6.9432 | -89.0442
6.9436 | -89.0465
6.9439 | -89.0485
6.9443 | -89.0507
6.9446 | -89.0536
6.9449 | -89.0566
6.9453 | -89.0589
6.9456 | -89.0603
6.946 | -89.062
6.9463 | -89.0639
6.9466 | -89.0662
6.947 | -89.0685
6.9473 | -89.0708
6.9477 | -89.0731
6.948 | -89.0755
6.9483 | -89.0776
6.9487 | -89.0796
6.949 | -89.082
6.9494 | -89.084
6.9497 | -89.0866
6.95 | -89.0892
6.9504 | -89.0914
6.9507 | -89.0942
6.9511 | -89.0972
6.9514 | -89.0998
6.9517 | -89.1019
6.9521 | -89.1037
6.9524 | -89.105
6.9528 | -89.1064
6.9531 | -89.1095
6.9534 | -89.1131
6.9538 | -89.1153
6.9541 | -89.1175
6.9545 | -89.1189
6.9548 | -89.1208
6.9552 | -89.1235
6.9555 | -89.1257
6.9558 | -89.1278
6.9562 | -89.1296
6.9565 | -89.1322
6.9569 | -89.1346
6.9572 | -89.1381
6.9575 | -89.1403
6.9579 | -89.1449
6.9582 | -89.1479
6.9586 | -89.1499
6.9589 | -89.1519
6.9592 | -89.1546
6.9596 | -89.1583
6.9599 | -89.1605
6.9603 | -89.1629
6.9606 | -89.1661
6.9609 | -89.1685
6.9613 | -89.1709
6.9616 | -89.1729
6.962 | -89.1751
6.9623 | -89.1769
6.9626 | -89.1793
6.963 | -89.1815
6.9633 | -89.1838
6.9637 | -89.1861
6.964 | -89.1883
6.9643 | -89.1904
6.9647 | -89.1926
6.965 | -89.1949
6.9654 | -89.1968
6.9657 | -89.1992
6.966 | -89.2012
6.9664 | -89.2038
6.9667 | -89.206
6.9671 | -89.2078
6.9674 | -89.2097
6.9677 | -89.212
6.9681 | -89.2157
6.9684 | -89.218
6.9688 | -89.2197
6.9691 | -89.2219
6.9695 | -89.2242
6.9698 | -89.2262
6.9701 | -89.228
6.9705 | -89.2298
6.9708 | -89.2319
6.9712 | -89.2338
6.9715 | -89.2357
6.9718 | -89.2381
6.9722 | -89.2408
6.9725 | -89.2433
6.9729 | -89.2455
6.9732 | -89.2475
6.9735 | -89.2495
6.9739 | -89.2516
6.9742 | -89.2542
6.9746 | -89.2568
6.9749 | -89.2594
6.9752 | -89.2616
6.9756 | -89.2627
6.9759 | -89.2643
6.9763 | -89.2653
6.9766 | -89.267
6.9769 | -89.2687
6.9773 | -89.2706
6.9776 | -89.2729
6.978 | -89.2764
6.9783 | -89.2785
6.9786 | -89.2814
6.979 | -89.2841
6.9793 | -89.2864
6.9797 | -89.2895
6.98 | -89.2932
6.9803 | -89.295
6.9807 | -89.2969
6.981 | -89.2991
6.9814 | -89.3013
6.9817 | -89.3035
6.982 | -89.3063
6.9824 | -89.3082
6.9827 | -89.3109
6.9831 | -89.3134
6.9834 | -89.3154
6.9838 | -89.3176
6.9841 | -89.3197
6.9844 | -89.3208
6.9848 | -89.3224
6.9851 | -89.3246
6.9855 | -89.3269
6.9858 | -89.3312
6.9861 | -89.3363
6.9865 | -89.3422
6.9868 | -89.3455
6.9872 | -89.3481
6.9875 | -89.3501
6.9878 | -89.3521
6.9882 | -89.354
6.9885 | -89.357
6.9889 | -89.3597
6.9892 | -89.3616
6.9895 | -89.3635
6.9899 | -89.3645
6.9902 | -89.3667
6.9906 | -89.3691
6.9909 | -89.3711
6.9912 | -89.3722
6.9916 | -89.3744
6.9919 | -89.379
6.9923 | -89.3815
6.9926 | -89.3836
6.9929 | -89.3855
6.9933 | -89.3876
6.9936 | -89.39
6.994 | -89.3923
6.9943 | -89.3946
6.9946 | -89.3968
6.995 | -89.3988
6.9953 | -89.4011
6.9957 | -89.4036
6.996 | -89.4054
6.9963 | -89.4061
6.9967 | -89.4068
6.997 | -89.4091
6.9974 | -89.411
6.9977 | -89.4126
6.9981 | -89.4143
6.9984 | -89.4147
6.9987 | -89.4152
6.9991 | -89.4171
6.9994 | -89.4189
6.9998 | -89.4202
7.0001 | -89.4218
7.0004 | -89.4236
7.0008 | -89.4251
7.0011 | -89.4259
7.0015 | -89.4281
7.0018 | -89.4314
7.0021 | -89.4332
7.0025 | -89.4354
7.0028 | -89.4372
7.0032 | -89.439
7.0035 | -89.4407
7.0038 | -89.4432
7.0042 | -89.4453
7.0045 | -89.4487
7.0049 | -89.4506
7.0052 | -89.4532
7.0055 | -89.4559
7.0059 | -89.4575
7.0062 | -89.4603
7.0066 | -89.4619
7.0069 | -89.4643
7.0072 | -89.4664
7.0076 | -89.4682
7.0079 | -89.469
7.0083 | -89.4718
7.0086 | -89.4755
7.0089 | -89.4781
7.0093 | -89.4798
7.0096 | -89.4816
7.01 | -89.4834
7.0103 | -89.4851
7.0107 | -89.487
7.011 | -89.4888
7.0113 | -89.49
7.0117 | -89.4922
7.012 | -89.4943
7.0124 | -89.4964
7.0127 | -89.4977
7.013 | -89.4992
7.0134 | -89.5014
7.0137 | -89.5036
7.0141 | -89.5048
7.0144 | -89.5061
7.0147 | -89.507
7.0151 | -89.5087
7.0154 | -89.5104
7.0158 | -89.5124
7.0161 | -89.5142
7.0164 | -89.5161
7.0168 | -89.5162
7.0171 | -89.5173
7.0175 | -89.5189
7.0178 | -89.5213
7.0181 | -89.5232
7.0185 | -89.5249
7.0188 | -89.528
7.0192 | -89.531
7.0195 | -89.5327
7.0198 | -89.5346
7.0202 | -89.5364
7.0205 | -89.5386
7.0209 | -89.5419
7.0212 | -89.544
7.0215 | -89.5464
7.0219 | -89.5495
7.0222 | -89.551
7.0226 | -89.5522
7.0229 | -89.5545
7.0232 | -89.5561
7.0236 | -89.5575
7.0239 | -89.559
7.0243 | -89.5604
7.0246 | -89.5623
7.025 | -89.564
7.0253 | -89.5658
7.0256 | -89.5673
7.026 | -89.5688
7.0263 | -89.5706
7.0267 | -89.5725
7.027 | -89.5742
7.0273 | -89.5752
7.0277 | -89.5769
7.028 | -89.5784
7.0284 | -89.5799
7.0287 | -89.5814
7.029 | -89.5831
7.0294 | -89.5859
7.0297 | -89.5881
7.0301 | -89.5895
7.0304 | -89.5901
7.0307 | -89.5912
7.0311 | -89.593
7.0314 | -89.5953
7.0318 | -89.5963
7.0321 | -89.5967
7.0324 | -89.5978
7.0328 | -89.5991
7.0331 | -89.6004
7.0335 | -89.6019
7.0338 | -89.6029
7.0341 | -89.6045
7.0345 | -89.6055
7.0348 | -89.6061
7.0352 | -89.6083
7.0355 | -89.61
7.0358 | -89.6122
7.0362 | -89.6144
7.0365 | -89.6157
7.0369 | -89.6172
7.0372 | -89.6188
7.0375 | -89.6202
7.0379 | -89.6226
7.0382 | -89.6248
7.0386 | -89.6268
7.0389 | -89.6288
7.0393 | -89.6303
7.0396 | -89.6328
7.0399 | -89.6346
7.0403 | -89.6346
7.0406 | -89.6346
7.041 | -89.6354
7.0413 | -89.6364
7.0416 | -89.6374
7.042 | -89.6394
7.0423 | -89.6409
7.0427 | -89.6421
7.043 | -89.6436
7.0433 | -89.6447
7.0437 | -89.6454
7.044 | -89.6472
7.0444 | -89.6487
7.0447 | -89.6502
7.045 | -89.6513
7.0454 | -89.6525
7.0457 | -89.6536
7.0461 | -89.6541
7.0464 | -89.6544
7.0467 | -89.6564
7.0471 | -89.658
7.0474 | -89.6594
7.0478 | -89.661
7.0481 | -89.6623
7.0484 | -89.6635
7.0488 | -89.6652
7.0491 | -89.6667
7.0495 | -89.6674
7.0498 | -89.668
7.0501 | -89.6698
7.0505 | -89.6711
7.0508 | -89.6729
7.0512 | -89.6738
7.0515 | -89.6752
7.0518 | -89.6761
7.0522 | -89.6775
7.0525 | -89.6776
7.0529 | -89.6783
7.0532 | -89.6792
7.0536 | -89.6801
7.0539 | -89.6815
7.0542 | -89.6829
7.0546 | -89.6841
7.0549 | -89.6852
7.0553 | -89.6861
7.0556 | -89.687
7.0559 | -89.6887
7.0563 | -89.69
7.0566 | -89.6903
7.057 | -89.69
7.0573 | -89.6894
7.0576 | -89.6898
7.058 | -89.6906
7.0583 | -89.6916
7.0587 | -89.6932
7.059 | -89.6944
7.0593 | -89.6955
7.0597 | -89.6965
7.06 | -89.6979
7.0604 | -89.6999
7.0607 | -89.7011
7.061 | -89.702
7.0614 | -89.7031
7.0617 | -89.7042
7.0621 | -89.7054
7.0624 | -89.7063
7.0627 | -89.7073
7.0631 | -89.708
7.0634 | -89.709
7.0638 | -89.7102
7.0641 | -89.7114
7.0644 | -89.7125
7.0648 | -89.7135
7.0651 | -89.715
7.0655 | -89.7162
7.0658 | -89.7172
7.0662 | -89.7182
7.0665 | -89.7194
7.0668 | -89.7206
7.0672 | -89.7229
7.0675 | -89.7248
7.0679 | -89.7262
7.0682 | -89.7275
7.0685 | -89.7286
7.0689 | -89.7293
7.0692 | -89.7308
7.0696 | -89.733
7.0699 | -89.7321
7.0702 | -89.7323
7.0706 | -89.7332
7.0709 | -89.7343
7.0713 | -89.7352
7.0716 | -89.7365
7.0719 | -89.738
7.0723 | -89.7392
7.0726 | -89.74
7.073 | -89.7407
7.0733 | -89.7411
7.0736 | -89.7408
7.074 | -89.7405
7.0743 | -89.7411
7.0747 | -89.7416
7.075 | -89.7422
7.0753 | -89.7437
7.0757 | -89.7455
7.076 | -89.7461
7.0764 | -89.7461
7.0767 | -89.7464
7.077 | -89.7471
7.0774 | -89.7475
7.0777 | -89.7479
7.0781 | -89.749
7.0784 | -89.7494
7.0787 | -89.7496
7.0791 | -89.7503
7.0794 | -89.7504
7.0798 | -89.7501
7.0801 | -89.75
7.0805 | -89.7503
7.0808 | -89.7507
7.0811 | -89.7509
7.0815 | -89.7514
7.0818 | -89.7513
7.0822 | -89.7525
7.0825 | -89.7534
7.0828 | -89.7537
7.0832 | -89.7541
7.0835 | -89.7546
7.0839 | -89.7549
7.0842 | -89.7571
7.0845 | -89.7593
7.0849 | -89.7596
7.0852 | -89.7596
7.0856 | -89.7605
7.0859 | -89.7632
7.0862 | -89.7638
7.0866 | -89.7648
7.0869 | -89.7663
7.0873 | -89.7674
7.0876 | -89.768
7.0879 | -89.7685
7.0883 | -89.7686
7.0886 | -89.7684
7.089 | -89.768
7.0893 | -89.7675
7.0896 | -89.7676
7.09 | -89.7678
7.0903 | -89.7685
7.0907 | -89.7689
7.091 | -89.7689
7.0913 | -89.7693
7.0917 | -89.7699
7.092 | -89.7699
7.0924 | -89.7703
7.0927 | -89.7689
7.093 | -89.7694
7.0934 | -89.7699
7.0937 | -89.7702
7.0941 | -89.7705
7.0944 | -89.7711
7.0948 | -89.7725
7.0951 | -89.775
7.0954 | -89.7761
7.0958 | -89.7765
7.0961 | -89.7774
7.0965 | -89.7782
7.0968 | -89.7798
7.0971 | -89.7801
7.0975 | -89.7802
7.0978 | -89.7801
7.0982 | -89.781
7.0985 | -89.7817
7.0988 | -89.783
7.0992 | -89.783
7.0995 | -89.7825
7.0999 | -89.7823
7.1002 | -89.782
7.1005 | -89.7816
7.1009 | -89.7818
7.1012 | -89.7822
7.1016 | -89.7828
7.1019 | -89.7833
7.1022 | -89.7828
7.1026 | -89.782
7.1029 | -89.7821
7.1033 | -89.7825
7.1036 | -89.7829
7.1039 | -89.7833
7.1043 | -89.7834
7.1046 | -89.7836
7.105 | -89.7841
7.1053 | -89.7841
7.1056 | -89.7844
7.106 | -89.784
7.1063 | -89.7842
7.1067 | -89.7841
7.107 | -89.7841
7.1073 | -89.7849
7.1077 | -89.7852
7.108 | -89.7844
7.1084 | -89.7846
7.1087 | -89.7839
7.1091 | -89.7843
7.1094 | -89.7854
7.1097 | -89.7864
7.1101 | -89.7873
7.1104 | -89.7874
7.1108 | -89.7862
7.1111 | -89.7859
7.1114 | -89.7859
7.1118 | -89.7861
7.1121 | -89.7868
7.1125 | -89.7865
7.1128 | -89.7862
7.1131 | -89.7858
7.1135 | -89.7862
7.1138 | -89.786
7.1142 | -89.7859
7.1145 | -89.7856
7.1148 | -89.7857
7.1152 | -89.7859
7.1155 | -89.7859
7.1159 | -89.7859
7.1162 | -89.7856
7.1165 | -89.7855
7.1169 | -89.7852
7.1172 | -89.7848
7.1176 | -89.7853
7.1179 | -89.7861
7.1182 | -89.7859
7.1186 | -89.7856
7.1189 | -89.7853
7.1193 | -89.785
7.1196 | -89.7845
7.1199 | -89.7845
7.1203 | -89.7838
7.1206 | -89.7839
7.121 | -89.7855
7.1213 | -89.7859
7.1216 | -89.7854
7.122 | -89.7855
7.1223 | -89.7845
7.1227 | -89.7847
7.123 | -89.7845
7.1234 | -89.7851
7.1237 | -89.7852
7.124 | -89.7851
7.1244 | -89.785
7.1247 | -89.7852
7.1251 | -89.785
7.1254 | -89.7853
7.1257 | -89.7853
7.1261 | -89.785
7.1264 | -89.785
7.1268 | -89.7848
7.1271 | -89.7844
7.1274 | -89.7835
7.1278 | -89.7819
7.1281 | -89.781
7.1285 | -89.7809
7.1288 | -89.7803
7.1291 | -89.7797
7.1295 | -89.7794
7.1298 | -89.7787
7.1302 | -89.7765
7.1305 | -89.776
7.1308 | -89.776
7.1312 | -89.7756
7.1315 | -89.7754
7.1319 | -89.7748
7.1322 | -89.7744
7.1325 | -89.7742
7.1329 | -89.7736
7.1332 | -89.773
7.1336 | -89.7725
7.1339 | -89.772
7.1342 | -89.7716
7.1346 | -89.7714
7.1349 | -89.7721
7.1353 | -89.7735
7.1356 | -89.7739
7.136 | -89.7738
7.1363 | -89.7734
7.1366 | -89.7744
7.137 | -89.7757
7.1373 | -89.7765
7.1377 | -89.777
7.138 | -89.7766
7.1383 | -89.7743
7.1387 | -89.7732
7.139 | -89.7724
7.1394 | -89.7726
7.1397 | -89.7736
7.14 | -89.7736
7.1404 | -89.773
7.1407 | -89.7726
7.1411 | -89.772
7.1414 | -89.7711
7.1417 | -89.7705
7.1421 | -89.7702
7.1424 | -89.7689
7.1428 | -89.7668
7.1431 | -89.767
7.1434 | -89.7669
7.1438 | -89.7656
7.1441 | -89.7646
7.1445 | -89.7637
7.1448 | -89.7629
7.1451 | -89.7628
7.1455 | -89.7631
7.1458 | -89.7623
7.1462 | -89.7606
7.1465 | -89.7588
7.1468 | -89.7582
7.1472 | -89.7579
7.1475 | -89.7578
7.1479 | -89.7568
7.1482 | -89.7563
7.1485 | -89.7558
7.1489 | -89.7554
7.1492 | -89.7552
7.1496 | -89.7537
7.1499 | -89.7526
7.1503 | -89.7513
7.1506 | -89.7504
7.1509 | -89.7497
7.1513 | -89.7486
7.1516 | -89.7468
7.152 | -89.7459
7.1523 | -89.7447
7.1526 | -89.7434
7.153 | -89.7426
7.1533 | -89.7413
7.1537 | -89.74
7.154 | -89.7385
7.1543 | -89.7376
7.1547 | -89.7368
7.155 | -89.7361
7.1554 | -89.7353
7.1557 | -89.7348
7.156 | -89.7342
7.1564 | -89.7339
7.1567 | -89.7326
7.1571 | -89.7314
7.1574 | -89.7315
7.1577 | -89.7312
7.1581 | -89.7304
7.1584 | -89.7297
7.1588 | -89.7298
7.1591 | -89.7295
7.1594 | -89.7286
7.1598 | -89.7278
7.1601 | -89.7272
7.1605 | -89.7266
7.1608 | -89.7285
7.1611 | -89.7297
7.1615 | -89.7276
7.1618 | -89.7267
7.1622 | -89.7257
7.1625 | -89.7253
7.1628 | -89.7245
7.1632 | -89.7236
7.1635 | -89.7225
7.1639 | -89.7205
7.1642 | -89.7174
7.1646 | -89.7154
7.1649 | -89.7141
7.1652 | -89.7138
7.1656 | -89.7129
7.1659 | -89.7109
7.1663 | -89.7098
7.1666 | -89.7093
7.1669 | -89.7088
7.1673 | -89.7073
7.1676 | -89.7054
7.168 | -89.7035
7.1683 | -89.7024
7.1686 | -89.7008
7.169 | -89.7002
7.1693 | -89.6988
7.1697 | -89.6978
7.17 | -89.6968
7.1703 | -89.6957
7.1707 | -89.6941
7.171 | -89.6932
7.1714 | -89.6922
7.1717 | -89.691
7.172 | -89.6903
7.1724 | -89.6889
7.1727 | -89.6876
7.1731 | -89.6876
7.1734 | -89.6858
7.1737 | -89.6825
7.1741 | -89.6806
7.1744 | -89.6802
7.1748 | -89.6793
7.1751 | -89.6784
7.1754 | -89.6773
7.1758 | -89.6763
7.1761 | -89.6755
7.1765 | -89.6746
7.1768 | -89.674
7.1771 | -89.6727
7.1775 | -89.6718
7.1778 | -89.6708
7.1782 | -89.6697
7.1785 | -89.6685
7.1789 | -89.6673
7.1792 | -89.667
7.1795 | -89.6667
7.1799 | -89.6659
7.1802 | -89.6649
7.1806 | -89.6638
7.1809 | -89.6629
7.1812 | -89.6614
7.1816 | -89.6605
7.1819 | -89.6594
7.1823 | -89.6579
7.1826 | -89.6566
7.1829 | -89.6557
7.1833 | -89.6554
7.1836 | -89.6543
7.184 | -89.6533
7.1843 | -89.652
7.1846 | -89.6508
7.185 | -89.6495
7.1853 | -89.6482
7.1857 | -89.6472
7.186 | -89.6468
7.1863 | -89.6473
7.1867 | -89.6462
7.187 | -89.6449
7.1874 | -89.6438
7.1877 | -89.6427
7.188 | -89.6416
7.1884 | -89.6405
7.1887 | -89.6397
7.1891 | -89.6386
7.1894 | -89.6372
7.1897 | -89.6345
7.1901 | -89.633
7.1904 | -89.6319
7.1908 | -89.631
7.1911 | -89.6304
7.1914 | -89.6297
7.1918 | -89.6282
7.1921 | -89.6271
7.1925 | -89.6263
7.1928 | -89.6255
7.1932 | -89.6248
7.1935 | -89.6246
7.1938 | -89.6238
7.1942 | -89.6227
7.1945 | -89.6218
7.1949 | -89.6207
7.1952 | -89.6208
7.1955 | -89.6205
7.1959 | -89.6195
7.1962 | -89.619
7.1966 | -89.6185
7.1969 | -89.6175
7.1972 | -89.6168
7.1976 | -89.6157
7.1979 | -89.6145
7.1983 | -89.6133
7.1986 | -89.6123
7.1989 | -89.6114
7.1993 | -89.6105
7.1996 | -89.6097
7.2 | -89.6092
7.2003 | -89.6086
7.2006 | -89.6077
7.201 | -89.6063
7.2013 | -89.6053
7.2017 | -89.6031
7.202 | -89.6024
7.2023 | -89.6021
7.2027 | -89.6015
7.203 | -89.6012
7.2034 | -89.6006
7.2037 | -89.5998
7.204 | -89.5985
7.2044 | -89.5971
7.2047 | -89.5956
7.2051 | -89.5946
7.2054 | -89.5932
7.2058 | -89.5917
7.2061 | -89.5904
7.2064 | -89.5892
7.2068 | -89.5884
7.2071 | -89.5877
7.2075 | -89.5859
7.2078 | -89.5847
7.2081 | -89.5836
7.2085 | -89.5827
7.2088 | -89.5819
7.2092 | -89.5809
7.2095 | -89.5798
7.2098 | -89.5786
7.2102 | -89.5776
7.2105 | -89.5771
7.2109 | -89.5759
7.2112 | -89.575
7.2115 | -89.5738
7.2119 | -89.5712
7.2122 | -89.5695
7.2126 | -89.5685
7.2129 | -89.567
7.2132 | -89.5659
7.2136 | -89.5647
7.2139 | -89.5636
7.2143 | -89.5629
7.2146 | -89.5628
7.2149 | -89.5621
7.2153 | -89.5612
7.2156 | -89.5609
7.216 | -89.5608
7.2163 | -89.56
7.2166 | -89.5583
7.217 | -89.5564
7.2173 | -89.5548
7.2177 | -89.5535
7.218 | -89.5526
7.2183 | -89.5518
7.2187 | -89.5511
7.219 | -89.5508
7.2194 | -89.5498
7.2197 | -89.5485
7.2201 | -89.5473
7.2204 | -89.5468
7.2207 | -89.5461
7.2211 | -89.5453
7.2214 | -89.5445
7.2218 | -89.5435
7.2221 | -89.5426
7.2224 | -89.5417
7.2228 | -89.5401
7.2231 | -89.5391
7.2235 | -89.5389
7.2238 | -89.538
7.2241 | -89.5367
7.2245 | -89.5354
7.2248 | -89.5344
7.2252 | -89.5332
7.2255 | -89.5322
7.2258 | -89.5309
7.2262 | -89.5296
7.2265 | -89.528
7.2269 | -89.5269
7.2272 | -89.5258
7.2275 | -89.5249
7.2279 | -89.5236
7.2282 | -89.5222
7.2286 | -89.5205
7.2289 | -89.5188
7.2292 | -89.5172
7.2296 | -89.5162
7.2299 | -89.5153
7.2303 | -89.5146
7.2306 | -89.5138
7.2309 | -89.5133
7.2313 | -89.5125
7.2316 | -89.5115
7.232 | -89.51
7.2323 | -89.509
7.2326 | -89.5077
7.233 | -89.506
7.2333 | -89.5048
7.2337 | -89.5036
7.234 | -89.5036
7.2344 | -89.5042
7.2347 | -89.5031
7.235 | -89.5021
7.2354 | -89.5012
7.2357 | -89.4999
7.2361 | -89.4987
7.2364 | -89.4974
7.2367 | -89.4959
7.2371 | -89.4948
7.2374 | -89.4941
7.2378 | -89.4933
7.2381 | -89.4921
7.2384 | -89.4906
7.2388 | -89.4886
7.2391 | -89.4874
7.2395 | -89.4868
7.2398 | -89.4855
7.2401 | -89.4839
7.2405 | -89.4822
7.2408 | -89.4807
7.2412 | -89.4795
7.2415 | -89.4786
7.2418 | -89.4778
7.2422 | -89.4768
7.2425 | -89.4758
7.2429 | -89.4747
7.2432 | -89.4741
7.2435 | -89.4732
7.2439 | -89.4722
7.2442 | -89.4695
7.2446 | -89.4671
7.2449 | -89.4664
7.2452 | -89.4662
7.2456 | -89.4651
7.2459 | -89.4638
7.2463 | -89.4625
7.2466 | -89.4616
7.2469 | -89.4609
7.2473 | -89.4602
7.2476 | -89.4599
7.248 | -89.4593
7.2483 | -89.459
7.2487 | -89.4582
7.249 | -89.4569
7.2493 | -89.4556
7.2497 | -89.4545
7.25 | -89.4543
7.2504 | -89.453
7.2507 | -89.4517
7.251 | -89.4507
7.2514 | -89.45
7.2517 | -89.4495
7.2521 | -89.4487
7.2524 | -89.4479
7.2527 | -89.4468
7.2531 | -89.4452
7.2534 | -89.4444
7.2538 | -89.4434
7.2541 | -89.4414
7.2544 | -89.44
7.2548 | -89.439
7.2551 | -89.438
7.2555 | -89.4363
7.2558 | -89.4352
7.2561 | -89.4344
7.2565 | -89.4334
7.2568 | -89.4325
7.2572 | -89.4316
7.2575 | -89.4303
7.2578 | -89.4287
7.2582 | -89.4279
7.2585 | -89.4269
7.2589 | -89.426
7.2592 | -89.4252
7.2595 | -89.4242
7.2599 | -89.4233
7.2602 | -89.4225
7.2606 | -89.4217
7.2609 | -89.4207
7.2613 | -89.4199
7.2616 | -89.4192
7.2619 | -89.4184
7.2623 | -89.4171
7.2626 | -89.4158
7.263 | -89.4148
7.2633 | -89.4141
7.2636 | -89.4132
7.264 | -89.4124
7.2643 | -89.4125
7.2647 | -89.4107
7.265 | -89.4091
7.2653 | -89.4084
7.2657 | -89.4077
7.266 | -89.4068
7.2664 | -89.4061
7.2667 | -89.4052
7.267 | -89.4043
7.2674 | -89.4035
7.2677 | -89.4027
7.2681 | -89.4029
7.2684 | -89.403
7.2687 | -89.4019
7.2691 | -89.4011
7.2694 | -89.4014
7.2698 | -89.4007
7.2701 | -89.4
7.2704 | -89.4003
7.2708 | -89.4005
7.2711 | -89.3997
7.2715 | -89.3987
7.2718 | -89.3979
7.2721 | -89.3972
7.2725 | -89.3972
7.2728 | -89.3968
7.2732 | -89.3959
7.2735 | -89.3952
7.2738 | -89.3947
7.2742 | -89.3932
7.2745 | -89.3918
7.2749 | -89.3906
7.2752 | -89.3897
7.2755 | -89.389
7.2759 | -89.3886
7.2762 | -89.3878
7.2766 | -89.3872
7.2769 | -89.3864
7.2773 | -89.3856
7.2776 | -89.3848
7.2779 | -89.3843
7.2783 | -89.3831
7.2786 | -89.3821
7.279 | -89.3813
7.2793 | -89.3817
7.2796 | -89.381
7.28 | -89.3804
7.2803 | -89.3797
7.2807 | -89.3792
7.281 | -89.3785
7.2813 | -89.3776
7.2817 | -89.3766
7.282 | -89.3756
7.2824 | -89.3745
7.2827 | -89.3729
7.283 | -89.3712
7.2834 | -89.3701
7.2837 | -89.3698
7.2841 | -89.3686
7.2844 | -89.3673
7.2847 | -89.3666
7.2851 | -89.3653
7.2854 | -89.3644
7.2858 | -89.3646
7.2861 | -89.3641
7.2864 | -89.3637
7.2868 | -89.3633
7.2871 | -89.3628
7.2875 | -89.3623
7.2878 | -89.3614
7.2881 | -89.3608
7.2885 | -89.3611
7.2888 | -89.3614
7.2892 | -89.3616
7.2895 | -89.3611
7.2899 | -89.3606
7.2902 | -89.36
7.2905 | -89.3596
7.2909 | -89.3592
7.2912 | -89.3587
7.2916 | -89.3589
7.2919 | -89.3593
7.2922 | -89.359
7.2926 | -89.3586
7.2929 | -89.3582
7.2933 | -89.3579
7.2936 | -89.3579
7.2939 | -89.3578
7.2943 | -89.3571
7.2946 | -89.3568
7.295 | -89.3562
7.2953 | -89.355
7.2956 | -89.3548
7.296 | -89.3545
7.2963 | -89.3538
7.2967 | -89.3532
7.297 | -89.3529
7.2973 | -89.3526
7.2977 | -89.3519
7.298 | -89.3515
7.2984 | -89.3512
7.2987 | -89.3507
7.299 | -89.3513
7.2994 | -89.351
7.2997 | -89.351
7.3001 | -89.3511
7.3004 | -89.3508
7.3007 | -89.3503
7.3011 | -89.35
7.3014 | -89.3496
7.3018 | -89.3493
7.3021 | -89.3486
7.3024 | -89.3486
7.3028 | -89.3481
7.3031 | -89.3471
7.3035 | -89.347
7.3038 | -89.3475
7.3042 | -89.3475
7.3045 | -89.3471
7.3048 | -89.3473
7.3052 | -89.3476
7.3055 | -89.3476
7.3059 | -89.3479
7.3062 | -89.349
7.3065 | -89.349
7.3069 | -89.3487
7.3072 | -89.3483
7.3076 | -89.3477
7.3079 | -89.3471
7.3082 | -89.3474
7.3086 | -89.3485
7.3089 | -89.3489
7.3093 | -89.3492
7.3096 | -89.3496
7.3099 | -89.3496
7.3103 | -89.349
7.3106 | -89.3485
7.311 | -89.348
7.3113 | -89.3478
7.3116 | -89.3478
7.312 | -89.347
7.3123 | -89.3461
7.3127 | -89.3458
7.313 | -89.3457
7.3133 | -89.3461
7.3137 | -89.3459
7.314 | -89.3457
7.3144 | -89.3453
7.3147 | -89.345
7.315 | -89.3455
7.3154 | -89.3458
7.3157 | -89.3456
7.3161 | -89.3452
7.3164 | -89.3453
7.3167 | -89.3448
7.3171 | -89.344
7.3174 | -89.3432
7.3178 | -89.343
7.3181 | -89.3429
7.3185 | -89.343
7.3188 | -89.3426
7.3191 | -89.3411
7.3195 | -89.3402
7.3198 | -89.34
7.3202 | -89.3397
7.3205 | -89.34
7.3208 | -89.3401
7.3212 | -89.3402
7.3215 | -89.3402
7.3219 | -89.3401
7.3222 | -89.3397
7.3225 | -89.3402
7.3229 | -89.3404
7.3232 | -89.34
7.3236 | -89.3392
7.3239 | -89.3391
7.3242 | -89.3389
7.3246 | -89.3391
7.3249 | -89.3394
7.3253 | -89.3393
7.3256 | -89.3391
7.3259 | -89.3394
7.3263 | -89.3402
7.3266 | -89.3404
7.327 | -89.3405
7.3273 | -89.3409
7.3276 | -89.3407
7.328 | -89.3405
7.3283 | -89.3416
7.3287 | -89.3418
7.329 | -89.342
7.3293 | -89.3417
7.3297 | -89.342
7.33 | -89.3429
7.3304 | -89.3431
7.3307 | -89.3433
7.3311 | -89.3434
7.3314 | -89.3434
7.3317 | -89.3434
7.3321 | -89.3437
7.3324 | -89.3441
7.3328 | -89.3442
7.3331 | -89.3454
7.3334 | -89.3464
7.3338 | -89.3477
7.3341 | -89.3484
7.3345 | -89.3486
7.3348 | -89.3488
7.3351 | -89.3486
7.3355 | -89.3483
7.3358 | -89.3485
7.3362 | -89.3487
7.3365 | -89.349
7.3368 | -89.3497
7.3372 | -89.3501
7.3375 | -89.3505
7.3379 | -89.3501
7.3382 | -89.3502
7.3385 | -89.3504
7.3389 | -89.3508
7.3392 | -89.3513
7.3396 | -89.3513
7.3399 | -89.3515
7.3402 | -89.3514
7.3406 | -89.3513
7.3409 | -89.3512
7.3413 | -89.3506
7.3416 | -89.3503
7.3419 | -89.3502
7.3423 | -89.35
7.3426 | -89.3499
7.343 | -89.3501
7.3433 | -89.3504
7.3436 | -89.3505
7.344 | -89.3506
7.3443 | -89.3521
7.3447 | -89.3531
7.345 | -89.3534
7.3453 | -89.3534
7.3457 | -89.3538
7.346 | -89.3535
7.3464 | -89.3526
7.3467 | -89.3524
7.3471 | -89.3532
7.3474 | -89.3537
7.3477 | -89.3543
7.3481 | -89.3551
7.3484 | -89.3555
7.3488 | -89.3555
7.3491 | -89.3556
7.3494 | -89.3559
7.3498 | -89.3563
7.3501 | -89.357
7.3505 | -89.3573
7.3508 | -89.3573
7.3511 | -89.3581
7.3515 | -89.3592
7.3518 | -89.3596
7.3522 | -89.36
7.3525 | -89.3606
7.3528 | -89.3617
7.3532 | -89.3622
7.3535 | -89.3627
7.3539 | -89.363
7.3542 | -89.3634
7.3545 | -89.3634
7.3549 | -89.3631
7.3552 | -89.3635
7.3556 | -89.364
7.3559 | -89.3642
7.3562 | -89.3641
7.3566 | -89.3647
7.3569 | -89.365
7.3573 | -89.3651
7.3576 | -89.3659
7.3579 | -89.3674
7.3583 | -89.3679
7.3586 | -89.3687
7.359 | -89.3692
7.3593 | -89.3699
7.3597 | -89.371
7.36 | -89.3717
7.3603 | -89.373
7.3607 | -89.3744
7.361 | -89.3749
7.3614 | -89.3753
7.3617 | -89.376
7.362 | -89.3761
7.3624 | -89.3764
7.3627 | -89.3769
7.3631 | -89.3774
7.3634 | -89.3774
7.3637 | -89.3782
7.3641 | -89.3791
7.3644 | -89.3789
7.3648 | -89.3793
7.3651 | -89.3799
7.3654 | -89.3802
7.3658 | -89.3806
7.3661 | -89.3816
7.3665 | -89.3821
7.3668 | -89.3826
7.3671 | -89.383
7.3675 | -89.3837
7.3678 | -89.385
7.3682 | -89.3855
7.3685 | -89.3861
7.3688 | -89.3873
7.3692 | -89.3883
7.3695 | -89.3889
7.3699 | -89.3895
7.3702 | -89.3898
7.3705 | -89.39
7.3709 | -89.3903
7.3712 | -89.3908
7.3716 | -89.3914
7.3719 | -89.392
7.3722 | -89.3932
7.3726 | -89.3937
7.3729 | -89.394
7.3733 | -89.3946
7.3736 | -89.3951
7.374 | -89.3956
7.3743 | -89.3965
7.3746 | -89.3972
7.375 | -89.3974
7.3753 | -89.3983
7.3757 | -89.3985
7.376 | -89.3988
7.3763 | -89.3996
7.3767 | -89.4003
7.377 | -89.4012
7.3774 | -89.402
7.3777 | -89.4031
7.378 | -89.4038
7.3784 | -89.4044
7.3787 | -89.405
7.3791 | -89.4056
7.3794 | -89.4062
7.3797 | -89.4073
7.3801 | -89.4082
7.3804 | -89.4086
7.3808 | -89.409
7.3811 | -89.4098
7.3814 | -89.4105
7.3818 | -89.411
7.3821 | -89.4112
7.3825 | -89.4118
7.3828 | -89.4124
7.3831 | -89.4136
7.3835 | -89.4148
7.3838 | -89.4152
7.3842 | -89.416
7.3845 | -89.4168
7.3848 | -89.4177
7.3852 | -89.4184
7.3855 | -89.4194
7.3859 | -89.4209
7.3862 | -89.422
7.3865 | -89.4229
7.3869 | -89.4237
7.3872 | -89.4248
7.3876 | -89.4264
7.3879 | -89.4275
7.3883 | -89.4277
7.3886 | -89.4286
7.3889 | -89.4297
7.3893 | -89.4305
7.3896 | -89.4313
7.39 | -89.4324
7.3903 | -89.4333
7.3906 | -89.434
7.391 | -89.435
7.3913 | -89.4363
7.3917 | -89.4374
7.392 | -89.4385
7.3923 | -89.4394
7.3927 | -89.4406
7.393 | -89.4421
7.3934 | -89.4437
7.3937 | -89.4451
7.394 | -89.4462
7.3944 | -89.4476
7.3947 | -89.4486
7.3951 | -89.4501
7.3954 | -89.4514
7.3957 | -89.4526
7.3961 | -89.4537
7.3964 | -89.4538
7.3968 | -89.4544
7.3971 | -89.4562
7.3974 | -89.4573
7.3978 | -89.4583
7.3981 | -89.4587
7.3985 | -89.4591
7.3988 | -89.46
7.3991 | -89.4613
7.3995 | -89.4624
7.3998 | -89.4636
7.4002 | -89.4646
7.4005 | -89.4653
7.4009 | -89.466
7.4012 | -89.4672
7.4015 | -89.4694
7.4019 | -89.4712
7.4022 | -89.4727
7.4026 | -89.4737
7.4029 | -89.4748
7.4032 | -89.4754
7.4036 | -89.4757
7.4039 | -89.4767
7.4043 | -89.4781
7.4046 | -89.4804
7.4049 | -89.4815
7.4053 | -89.4821
7.4056 | -89.4833
7.406 | -89.4843
7.4063 | -89.486
7.4066 | -89.4876
7.407 | -89.4892
7.4073 | -89.4915
7.4077 | -89.4932
7.408 | -89.4939
7.4083 | -89.495
7.4087 | -89.4961
7.409 | -89.4972
7.4094 | -89.4983
7.4097 | -89.4994
7.41 | -89.5009
7.4104 | -89.5019
7.4107 | -89.5029
7.4111 | -89.5044
7.4114 | -89.5054
7.4117 | -89.5059
7.4121 | -89.5078
7.4124 | -89.509
7.4128 | -89.5105
7.4131 | -89.5117
7.4134 | -89.5128
7.4138 | -89.5139
7.4141 | -89.515
7.4145 | -89.5169
7.4148 | -89.5182
7.4151 | -89.5193
7.4155 | -89.5209
7.4158 | -89.5226
7.4162 | -89.523
7.4165 | -89.5242
7.4169 | -89.5259
7.4172 | -89.5272
7.4175 | -89.5282
7.4179 | -89.5297
7.4182 | -89.5303
7.4186 | -89.5313
7.4189 | -89.5324
7.4192 | -89.5327
7.4196 | -89.5329
7.4199 | -89.5336
7.4203 | -89.5346
7.4206 | -89.5354
7.4209 | -89.5363
7.4213 | -89.5375
7.4216 | -89.5387
7.422 | -89.5398
7.4223 | -89.5411
7.4226 | -89.5424
7.423 | -89.5433
7.4233 | -89.5447
7.4237 | -89.546
7.424 | -89.5471
7.4243 | -89.5484
7.4247 | -89.5501
7.425 | -89.5512
7.4254 | -89.5519
7.4257 | -89.553
7.426 | -89.5543
7.4264 | -89.5556
7.4267 | -89.5577
7.4271 | -89.559
7.4274 | -89.5602
7.4277 | -89.5616
7.4281 | -89.5626
7.4284 | -89.5638
7.4288 | -89.5656
7.4291 | -89.5669
7.4295 | -89.5678
7.4298 | -89.5692
7.4301 | -89.5705
7.4305 | -89.572
7.4308 | -89.5732
7.4312 | -89.5741
7.4315 | -89.5753
7.4318 | -89.5759
7.4322 | -89.577
7.4325 | -89.5788
7.4329 | -89.5806
7.4332 | -89.5819
7.4335 | -89.5835
7.4339 | -89.5857
7.4342 | -89.5869
7.4346 | -89.5883
7.4349 | -89.5898
7.4352 | -89.5914
7.4356 | -89.593
7.4359 | -89.5944
7.4363 | -89.5954
7.4366 | -89.5974
7.4369 | -89.5986
7.4373 | -89.6
7.4376 | -89.6014
7.438 | -89.6029
7.4383 | -89.6043
7.4386 | -89.6053
7.439 | -89.6067
7.4393 | -89.6086
7.4397 | -89.6102
7.44 | -89.6116
7.4403 | -89.6129
7.4407 | -89.6144
7.441 | -89.6166
7.4414 | -89.6185
7.4417 | -89.62
7.442 | -89.6211
7.4424 | -89.6221
7.4427 | -89.6235
7.4431 | -89.6252
7.4434 | -89.6265
7.4438 | -89.6274
7.4441 | -89.6277
7.4444 | -89.6287
7.4448 | -89.6306
7.4451 | -89.6328
7.4455 | -89.6347
7.4458 | -89.6364
7.4461 | -89.6378
7.4465 | -89.6391
7.4468 | -89.6406
7.4472 | -89.642
7.4475 | -89.6435
7.4478 | -89.6451
7.4482 | -89.6467
7.4485 | -89.6483
7.4489 | -89.6484
7.4492 | -89.6495
7.4495 | -89.6506
7.4499 | -89.6523
7.4502 | -89.6546
7.4506 | -89.6565
7.4509 | -89.6581
7.4512 | -89.6588
7.4516 | -89.6603
7.4519 | -89.6625
7.4523 | -89.665
7.4526 | -89.6673
7.4529 | -89.6687
7.4533 | -89.6702
7.4536 | -89.6717
7.454 | -89.6736
7.4543 | -89.6755
7.4546 | -89.6773
7.455 | -89.6788
7.4553 | -89.6804
7.4557 | -89.6818
7.456 | -89.6831
7.4563 | -89.6853
7.4567 | -89.6873
7.457 | -89.6888
7.4574 | -89.6905
7.4577 | -89.6919
7.4581 | -89.6933
7.4584 | -89.6949
7.4587 | -89.6967
7.4591 | -89.6989
7.4594 | -89.7009
7.4598 | -89.7021
7.4601 | -89.7036
7.4604 | -89.7049
7.4608 | -89.7062
7.4611 | -89.7075
7.4615 | -89.7094
7.4618 | -89.7107
7.4621 | -89.7115
7.4625 | -89.7132
7.4628 | -89.7155
7.4632 | -89.7174
7.4635 | -89.7193
7.4638 | -89.7211
7.4642 | -89.7224
7.4645 | -89.7236
7.4649 | -89.7253
7.4652 | -89.7271
7.4655 | -89.7289
7.4659 | -89.731
7.4662 | -89.7325
7.4666 | -89.7339
7.4669 | -89.7355
7.4672 | -89.7373
7.4676 | -89.7388
7.4679 | -89.7407
7.4683 | -89.7419
7.4686 | -89.7433
7.4689 | -89.7447
7.4693 | -89.7463
7.4696 | -89.7484
7.47 | -89.7507
7.4703 | -89.7524
7.4707 | -89.7538
7.471 | -89.7553
7.4713 | -89.7569
7.4717 | -89.7586
7.472 | -89.7605
7.4724 | -89.7628
7.4727 | -89.7649
7.473 | -89.7662
7.4734 | -89.7678
7.4737 | -89.7692
7.4741 | -89.7702
7.4744 | -89.7719
7.4747 | -89.7737
7.4751 | -89.7753
7.4754 | -89.7766
7.4758 | -89.7787
7.4761 | -89.7803
7.4764 | -89.782
7.4768 | -89.7836
7.4771 | -89.7856
7.4775 | -89.7886
7.4778 | -89.7909
7.4781 | -89.794
7.4785 | -89.7964
7.4788 | -89.7988
7.4792 | -89.8007
7.4795 | -89.8024
7.4798 | -89.804
7.4802 | -89.8057
7.4805 | -89.8077
7.4809 | -89.8099
7.4812 | -89.8118
7.4815 | -89.8136
7.4819 | -89.8159
7.4822 | -89.8184
7.4826 | -89.8204
7.4829 | -89.8222
7.4832 | -89.8241
7.4836 | -89.826
7.4839 | -89.8277
7.4843 | -89.8294
7.4846 | -89.8312
7.485 | -89.8331
7.4853 | -89.835
7.4856 | -89.8373
7.486 | -89.8391
7.4863 | -89.8406
7.4867 | -89.843
7.487 | -89.8455
7.4873 | -89.8475
7.4877 | -89.8492
7.488 | -89.8508
7.4884 | -89.8527
7.4887 | -89.8546
7.489 | -89.8568
7.4894 | -89.8594
7.4897 | -89.8623
7.4901 | -89.8643
7.4904 | -89.8661
7.4907 | -89.868
7.4911 | -89.8694
7.4914 | -89.8706
7.4918 | -89.8721
7.4921 | -89.8741
7.4924 | -89.877
7.4928 | -89.8798
7.4931 | -89.8816
7.4935 | -89.8835
7.4938 | -89.8862
7.4941 | -89.8883
7.4945 | -89.89
7.4948 | -89.8918
7.4952 | -89.893
7.4955 | -89.8947
7.4958 | -89.8963
7.4962 | -89.8979
7.4965 | -89.8994
7.4969 | -89.9018
7.4972 | -89.9035
7.4975 | -89.9053
7.4979 | -89.9067
7.4982 | -89.9083
7.4986 | -89.9105
7.4989 | -89.9123
7.4993 | -89.9142
7.4996 | -89.9173
7.4999 | -89.921
7.5003 | -89.923
7.5006 | -89.9251
7.501 | -89.9271
7.5013 | -89.9289
7.5016 | -89.9325
7.502 | -89.9364
7.5023 | -89.9382
7.5027 | -89.9403
7.503 | -89.9427
7.5033 | -89.9447
7.5037 | -89.9466
7.504 | -89.9488
7.5044 | -89.9517
7.5047 | -89.9539
7.505 | -89.9557
7.5054 | -89.958
7.5057 | -89.9595
7.5061 | -89.961
7.5064 | -89.9629
7.5067 | -89.9648
7.5071 | -89.9668
7.5074 | -89.9687
7.5078 | -89.9703
7.5081 | -89.9708
7.5084 | -89.9715
7.5088 | -89.973
7.5091 | -89.9751
7.5095 | -89.9778
7.5098 | -89.9797
7.5101 | -89.9815
7.5105 | -89.9833
7.5108 | -89.985
7.5112 | -89.9868
7.5115 | -89.9886
7.5118 | -89.9917
7.5122 | -89.9938
7.5125 | -89.9951
7.5129 | -89.9961
7.5132 | -89.998
7.5136 | -89.9994
7.5139 | -90.0014
7.5142 | -90.0036
7.5146 | -90.0054
7.5149 | -90.0077
7.5153 | -90.0101
7.5156 | -90.0126
7.5159 | -90.0144
7.5163 | -90.0164
7.5166 | -90.0181
7.517 | -90.0205
7.5173 | -90.0225
7.5176 | -90.0237
7.518 | -90.0252
7.5183 | -90.0271
7.5187 | -90.0296
7.519 | -90.032
7.5193 | -90.0335
7.5197 | -90.035
7.52 | -90.0372
7.5204 | -90.0389
7.5207 | -90.0405
7.521 | -90.0428
7.5214 | -90.0468
7.5217 | -90.05
7.5221 | -90.0513
7.5224 | -90.0524
7.5227 | -90.0541
7.5231 | -90.0559
7.5234 | -90.0567
7.5238 | -90.0584
7.5241 | -90.0603
7.5244 | -90.0619
7.5248 | -90.0628
7.5251 | -90.0642
7.5255 | -90.0661
7.5258 | -90.0683
7.5261 | -90.07
7.5265 | -90.0716
7.5268 | -90.0736
7.5272 | -90.0758
7.5275 | -90.0779
7.5279 | -90.0792
7.5282 | -90.0811
7.5285 | -90.0828
7.5289 | -90.084
7.5292 | -90.0861
7.5296 | -90.0877
7.5299 | -90.0891
7.5302 | -90.0912
7.5306 | -90.0925
7.5309 | -90.0942
7.5313 | -90.0961
7.5316 | -90.098
7.5319 | -90.0998
7.5323 | -90.1012
7.5326 | -90.1035
7.533 | -90.1059
7.5333 | -90.1078
7.5336 | -90.1098
7.534 | -90.1115
7.5343 | -90.113
7.5347 | -90.1147
7.535 | -90.1168
7.5353 | -90.1186
7.5357 | -90.1198
7.536 | -90.1219
7.5364 | -90.1244
7.5367 | -90.1263
7.537 | -90.127
7.5374 | -90.1295
7.5377 | -90.1321
7.5381 | -90.134
7.5384 | -90.1361
7.5387 | -90.1375
7.5391 | -90.139
7.5394 | -90.1411
7.5398 | -90.1433
7.5401 | -90.1449
7.5404 | -90.1475
7.5408 | -90.1502
7.5411 | -90.152
7.5415 | -90.1538
7.5418 | -90.1557
7.5422 | -90.1574
7.5425 | -90.1595
7.5428 | -90.1616
7.5432 | -90.1633
7.5435 | -90.1649
7.5439 | -90.1665
7.5442 | -90.1684
7.5445 | -90.1707
7.5449 | -90.1727
7.5452 | -90.1745
7.5456 | -90.1766
7.5459 | -90.1788
7.5462 | -90.1813
7.5466 | -90.1848
7.5469 | -90.1882
7.5473 | -90.1905
7.5476 | -90.1925
7.5479 | -90.195
7.5483 | -90.197
7.5486 | -90.1987
7.549 | -90.2006
7.5493 | -90.2024
7.5496 | -90.2041
7.55 | -90.2059
7.5503 | -90.2071
7.5507 | -90.2082
7.551 | -90.2099
7.5513 | -90.2119
7.5517 | -90.2143
7.552 | -90.2164
7.5524 | -90.2183
7.5527 | -90.2211
7.553 | -90.2242
7.5534 | -90.2259
7.5537 | -90.228
7.5541 | -90.2299
7.5544 | -90.2309
7.5548 | -90.2324
7.5551 | -90.2341
7.5554 | -90.2362
7.5558 | -90.2381
7.5561 | -90.24
7.5565 | -90.2427
7.5568 | -90.2452
7.5571 | -90.2472
7.5575 | -90.249
7.5578 | -90.251
7.5582 | -90.2532
7.5585 | -90.2551
7.5588 | -90.2574
7.5592 | -90.2591
7.5595 | -90.262
7.5599 | -90.2648
7.5602 | -90.2668
7.5605 | -90.268
7.5609 | -90.2697
7.5612 | -90.2719
7.5616 | -90.274
7.5619 | -90.2757
7.5622 | -90.2775
7.5626 | -90.2796
7.5629 | -90.2818
7.5633 | -90.2838
7.5636 | -90.2859
7.5639 | -90.2884
7.5643 | -90.2893
7.5646 | -90.2926
7.565 | -90.2965
7.5653 | -90.299
7.5656 | -90.3011
7.566 | -90.3032
7.5663 | -90.3061
7.5667 | -90.3087
7.567 | -90.3109
7.5673 | -90.3128
7.5677 | -90.3145
7.568 | -90.3157
7.5684 | -90.3172
7.5687 | -90.3191
7.5691 | -90.3207
7.5694 | -90.3226
7.5697 | -90.3251
7.5701 | -90.3271
7.5704 | -90.329
7.5708 | -90.3303
7.5711 | -90.3308
7.5714 | -90.3326
7.5718 | -90.3342
7.5721 | -90.3357
7.5725 | -90.3372
7.5728 | -90.339
7.5731 | -90.3411
7.5735 | -90.3425
7.5738 | -90.3434
7.5742 | -90.3442
7.5745 | -90.346
7.5748 | -90.3475
7.5752 | -90.3487
7.5755 | -90.3505
7.5759 | -90.3527
7.5762 | -90.3543
7.5765 | -90.3557
7.5769 | -90.3575
7.5772 | -90.3593
7.5776 | -90.3611
7.5779 | -90.363
7.5782 | -90.3666
7.5786 | -90.3688
7.5789 | -90.3701
7.5793 | -90.3717
7.5796 | -90.3734
7.5799 | -90.375
7.5803 | -90.377
7.5806 | -90.3787
7.581 | -90.3805
7.5813 | -90.3822
7.5816 | -90.384
7.582 | -90.3858
7.5823 | -90.3884
7.5827 | -90.3917
7.583 | -90.3944
7.5834 | -90.3965
7.5837 | -90.3984
7.584 | -90.3999
7.5844 | -90.4015
7.5847 | -90.4033
7.5851 | -90.4054
7.5854 | -90.4079
7.5857 | -90.4099
7.5861 | -90.4116
7.5864 | -90.4136
7.5868 | -90.4156
7.5871 | -90.4171
7.5874 | -90.4188
7.5878 | -90.4209
7.5881 | -90.4217
7.5885 | -90.4237
7.5888 | -90.4263
7.5891 | -90.4283
7.5895 | -90.43
7.5898 | -90.432
7.5902 | -90.4346
7.5905 | -90.4362
7.5908 | -90.4383
7.5912 | -90.4396
7.5915 | -90.4405
7.5919 | -90.4422
7.5922 | -90.4438
7.5925 | -90.4454
7.5929 | -90.4466
7.5932 | -90.4483
7.5936 | -90.4502
7.5939 | -90.4518
7.5942 | -90.4536
7.5946 | -90.456
7.5949 | -90.4582
7.5953 | -90.4601
7.5956 | -90.462
7.596 | -90.4634
7.5963 | -90.465
7.5966 | -90.4666
7.597 | -90.4677
7.5973 | -90.4688
7.5977 | -90.4703
7.598 | -90.4723
7.5983 | -90.4745
7.5987 | -90.4769
7.599 | -90.4798
7.5994 | -90.4814
7.5997 | -90.4827
7.6 | -90.4839
7.6004 | -90.4858
7.6007 | -90.4875
7.6011 | -90.4888
7.6014 | -90.4898
7.6017 | -90.491
7.6021 | -90.4918
7.6024 | -90.4938
7.6028 | -90.4961
7.6031 | -90.4973
7.6034 | -90.4985
7.6038 | -90.4994
7.6041 | -90.5009
7.6045 | -90.5024
7.6048 | -90.504
7.6051 | -90.5052
7.6055 | -90.5062
7.6058 | -90.5066
7.6062 | -90.5078
7.6065 | -90.5093
7.6068 | -90.511
7.6072 | -90.5131
7.6075 | -90.515
7.6079 | -90.5168
7.6082 | -90.5185
7.6085 | -90.5197
7.6089 | -90.5209
7.6092 | -90.5219
7.6096 | -90.5238
7.6099 | -90.5263
7.6102 | -90.5283
7.6106 | -90.5303
7.6109 | -90.5322
7.6113 | -90.5337
7.6116 | -90.5349
7.612 | -90.5361
7.6123 | -90.5376
7.6126 | -90.5397
7.613 | -90.542
7.6133 | -90.5424
7.6137 | -90.5434
7.614 | -90.5447
7.6143 | -90.5462
7.6147 | -90.5474
7.615 | -90.5485
7.6154 | -90.5502
7.6157 | -90.5517
7.616 | -90.5534
7.6164 | -90.5549
7.6167 | -90.5561
7.6171 | -90.5579
7.6174 | -90.5591
7.6177 | -90.5606
7.6181 | -90.5619
7.6184 | -90.5632
7.6188 | -90.5645
7.6191 | -90.5669
7.6194 | -90.5684
7.6198 | -90.5708
7.6201 | -90.573
7.6205 | -90.5745
7.6208 | -90.5762
7.6211 | -90.5771
7.6215 | -90.5783
7.6218 | -90.5808
7.6222 | -90.5827
7.6225 | -90.5841
7.6228 | -90.5854
7.6232 | -90.5869
7.6235 | -90.5884
7.6239 | -90.5898
7.6242 | -90.5911
7.6246 | -90.5932
7.6249 | -90.5953
7.6252 | -90.5968
7.6256 | -90.5981
7.6259 | -90.5996
7.6263 | -90.6014
7.6266 | -90.6029
7.6269 | -90.6044
7.6273 | -90.6062
7.6276 | -90.6083
7.628 | -90.61
7.6283 | -90.6113
7.6286 | -90.6129
7.629 | -90.6151
7.6293 | -90.617
7.6297 | -90.6178
7.63 | -90.6192
7.6303 | -90.6208
7.6307 | -90.6232
7.631 | -90.6253
7.6314 | -90.6273
7.6317 | -90.6291
7.632 | -90.6318
7.6324 | -90.6338
7.6327 | -90.6346
7.6331 | -90.6359
7.6334 | -90.6374
7.6337 | -90.6389
7.6341 | -90.6398
7.6344 | -90.6414
7.6348 | -90.6429
7.6351 | -90.6439
7.6354 | -90.6453
7.6358 | -90.6466
7.6361 | -90.648
7.6365 | -90.6488
7.6368 | -90.6487
7.6371 | -90.6487
7.6375 | -90.6496
7.6378 | -90.6508
7.6382 | -90.6519
7.6385 | -90.6533
7.6389 | -90.6544
7.6392 | -90.6552
7.6395 | -90.6562
7.6399 | -90.6583
7.6402 | -90.6609
7.6406 | -90.6628
7.6409 | -90.6644
7.6412 | -90.6662
7.6416 | -90.6683
7.6419 | -90.6699
7.6423 | -90.6715
7.6426 | -90.6746
7.6429 | -90.6748
7.6433 | -90.6766
7.6436 | -90.6793
7.644 | -90.681
7.6443 | -90.6823
7.6446 | -90.6836
7.645 | -90.6849
7.6453 | -90.6862
7.6457 | -90.6872
7.646 | -90.6883
7.6463 | -90.6901
7.6467 | -90.6915
7.647 | -90.6927
7.6474 | -90.6941
7.6477 | -90.6953
7.648 | -90.6964
7.6484 | -90.6976
7.6487 | -90.6985
7.6491 | -90.6992
7.6494 | -90.6999
7.6497 | -90.7011
7.6501 | -90.7025
7.6504 | -90.7039
7.6508 | -90.7052
7.6511 | -90.7068
7.6514 | -90.71
7.6518 | -90.7118
7.6521 | -90.7132
7.6525 | -90.7145
7.6528 | -90.7154
7.6532 | -90.7162
7.6535 | -90.7172
7.6538 | -90.7184
7.6542 | -90.7194
7.6545 | -90.7192
7.6549 | -90.7193
7.6552 | -90.7203
7.6555 | -90.7216
7.6559 | -90.7228
7.6562 | -90.725
7.6566 | -90.7278
7.6569 | -90.729
7.6572 | -90.7305
7.6576 | -90.7324
7.6579 | -90.7341
7.6583 | -90.7354
7.6586 | -90.7365
7.6589 | -90.7374
7.6593 | -90.7391
7.6596 | -90.7409
7.66 | -90.7436
7.6603 | -90.7457
7.6606 | -90.7471
7.661 | -90.7483
7.6613 | -90.7497
7.6617 | -90.7514
7.662 | -90.7528
7.6623 | -90.7541
7.6627 | -90.7553
7.663 | -90.7565
7.6634 | -90.7579
7.6637 | -90.7596
7.664 | -90.7613
7.6644 | -90.7632
7.6647 | -90.7642
7.6651 | -90.7651
7.6654 | -90.7669
7.6658 | -90.7686
7.6661 | -90.7699
7.6664 | -90.7712
7.6668 | -90.7725
7.6671 | -90.7743
7.6675 | -90.7758
7.6678 | -90.7775
7.6681 | -90.7792
7.6685 | -90.7808
7.6688 | -90.7815
7.6692 | -90.7819
7.6695 | -90.7826
7.6698 | -90.7839
7.6702 | -90.7849
7.6705 | -90.786
7.6709 | -90.787
7.6712 | -90.7888
7.6715 | -90.7905
7.6719 | -90.7916
7.6722 | -90.7921
7.6726 | -90.7932
7.6729 | -90.7943
7.6732 | -90.7952
7.6736 | -90.7962
7.6739 | -90.7969
7.6743 | -90.7976
7.6746 | -90.7986
7.6749 | -90.7996
7.6753 | -90.8006
7.6756 | -90.8015
7.676 | -90.8026
7.6763 | -90.8039
7.6766 | -90.8047
7.677 | -90.8061
7.6773 | -90.8075
7.6777 | -90.8087
7.678 | -90.8096
7.6783 | -90.8107
7.6787 | -90.8119
7.679 | -90.8145
7.6794 | -90.8181
7.6797 | -90.8204
7.68 | -90.8218
7.6804 | -90.8223
7.6807 | -90.8233
7.6811 | -90.8254
7.6814 | -90.8268
7.6818 | -90.8277
7.6821 | -90.8289
7.6824 | -90.8301
7.6828 | -90.8317
7.6831 | -90.8339
7.6835 | -90.8361
7.6838 | -90.8372
7.6841 | -90.8388
7.6845 | -90.8391
7.6848 | -90.8402
7.6852 | -90.841
7.6855 | -90.8418
7.6858 | -90.8428
7.6862 | -90.8445
7.6865 | -90.8455
7.6869 | -90.8465
7.6872 | -90.8473
7.6875 | -90.8481
7.6879 | -90.8485
7.6882 | -90.8492
7.6886 | -90.8501
7.6889 | -90.8512
7.6892 | -90.854
7.6896 | -90.8576
7.6899 | -90.8584
7.6903 | -90.8587
7.6906 | -90.8598
7.6909 | -90.8607
7.6913 | -90.8622
7.6916 | -90.8632
7.692 | -90.8643
7.6923 | -90.8653
7.6926 | -90.8661
7.693 | -90.8661
7.6933 | -90.8669
7.6937 | -90.8677
7.694 | -90.8684
7.6944 | -90.8691
7.6947 | -90.8701
7.695 | -90.8717
7.6954 | -90.8725
7.6957 | -90.8734
7.6961 | -90.8742
7.6964 | -90.8749
7.6967 | -90.8756
7.6971 | -90.8763
7.6974 | -90.8777
7.6978 | -90.8787
7.6981 | -90.8794
7.6984 | -90.8804
7.6988 | -90.8816
7.6991 | -90.8833
7.6995 | -90.8845
7.6998 | -90.8852
7.7001 | -90.886
7.7005 | -90.8866
7.7008 | -90.8877
7.7012 | -90.8886
7.7015 | -90.8896
7.7018 | -90.8901
7.7022 | -90.8916
7.7025 | -90.893
7.7029 | -90.8936
7.7032 | -90.8937
7.7035 | -90.8942
7.7039 | -90.8946
7.7042 | -90.8948
7.7046 | -90.8954
7.7049 | -90.8962
7.7052 | -90.8972
7.7056 | -90.898
7.7059 | -90.8988
7.7063 | -90.9
7.7066 | -90.9025
7.7069 | -90.9031
7.7073 | -90.9033
7.7076 | -90.9038
7.708 | -90.9047
7.7083 | -90.9054
7.7087 | -90.9064
7.709 | -90.907
7.7093 | -90.9072
7.7097 | -90.9073
7.71 | -90.9076
7.7104 | -90.9087
7.7107 | -90.9099
7.711 | -90.9107
7.7114 | -90.9108
7.7117 | -90.9113
7.7121 | -90.9122
7.7124 | -90.9136
7.7127 | -90.9144
7.7131 | -90.9152
7.7134 | -90.9155
7.7138 | -90.9162
7.7141 | -90.9169
7.7144 | -90.9173
7.7148 | -90.9175
7.7151 | -90.918
7.7155 | -90.9191
7.7158 | -90.9201
7.7161 | -90.9206
7.7165 | -90.9209
7.7168 | -90.9216
7.7172 | -90.922
7.7175 | -90.9228
7.7178 | -90.9235
7.7182 | -90.9242
7.7185 | -90.9251
7.7189 | -90.9267
7.7192 | -90.9285
7.7195 | -90.9304
7.7199 | -90.9316
7.7202 | -90.9322
7.7206 | -90.933
7.7209 | -90.9338
7.7212 | -90.9347
7.7216 | -90.9358
7.7219 | -90.9368
7.7223 | -90.9379
7.7226 | -90.9389
7.723 | -90.9397
7.7233 | -90.9403
7.7236 | -90.9408
7.724 | -90.942
7.7243 | -90.9431
7.7247 | -90.9436
7.725 | -90.9439
7.7253 | -90.9452
7.7257 | -90.9474
7.726 | -90.9488
7.7264 | -90.9495
7.7267 | -90.9502
7.727 | -90.9515
7.7274 | -90.9524
7.7277 | -90.9529
7.7281 | -90.9534
7.7284 | -90.9538
7.7287 | -90.9539
7.7291 | -90.954
7.7294 | -90.9545
7.7298 | -90.9551
7.7301 | -90.9559
7.7304 | -90.9588
7.7308 | -90.9602
7.7311 | -90.9609
7.7315 | -90.9616
7.7318 | -90.9624
7.7321 | -90.9627
7.7325 | -90.9628
7.7328 | -90.9635
7.7332 | -90.9649
7.7335 | -90.9661
7.7338 | -90.9663
7.7342 | -90.9663
7.7345 | -90.9662
7.7349 | -90.9669
7.7352 | -90.9687
7.7355 | -90.9702
7.7359 | -90.9717
7.7362 | -90.9732
7.7366 | -90.9744
7.7369 | -90.975
7.7373 | -90.9756
7.7376 | -90.9763
7.7379 | -90.9764
7.7383 | -90.9762
7.7386 | -90.977
7.739 | -90.978
7.7393 | -90.9786
7.7396 | -90.9789
7.74 | -90.979
7.7403 | -90.9811
7.7407 | -90.9825
7.741 | -90.9825
7.7413 | -90.9831
7.7417 | -90.9841
7.742 | -90.9846
7.7424 | -90.9857
7.7427 | -90.9862
7.743 | -90.9862
7.7434 | -90.9868
7.7437 | -90.9877
7.7441 | -90.9882
7.7444 | -90.9888
7.7447 | -90.9894
7.7451 | -90.9897
7.7454 | -90.9904
7.7458 | -90.9909
7.7461 | -90.9917
7.7464 | -90.9922
7.7468 | -90.9927
7.7471 | -90.993
7.7475 | -90.993
7.7478 | -90.994
7.7481 | -90.9955
7.7485 | -90.9968
7.7488 | -90.9979
7.7492 | -90.9983
7.7495 | -90.9986
7.7499 | -90.9993
7.7502 | -90.9998
7.7505 | -90.9999
7.7509 | -91.0005
7.7512 | -91.001
7.7516 | -91.0009
7.7519 | -91.0013
7.7522 | -91.0024
7.7526 | -91.0027
7.7529 | -91.0033
7.7533 | -91.0046
7.7536 | -91.0052
7.7539 | -91.0058
7.7543 | -91.0063
7.7546 | -91.0067
7.755 | -91.0075
7.7553 | -91.0087
7.7556 | -91.0094
7.756 | -91.0102
7.7563 | -91.0107
7.7567 | -91.0112
7.757 | -91.0122
7.7573 | -91.0128
7.7577 | -91.0133
7.758 | -91.0137
7.7584 | -91.0143
7.7587 | -91.0155
7.759 | -91.017
7.7594 | -91.0178
7.7597 | -91.0182
7.7601 | -91.0192
7.7604 | -91.0206
7.7607 | -91.0212
7.7611 | -91.0224
7.7614 | -91.0236
7.7618 | -91.0246
7.7621 | -91.0252
7.7624 | -91.0257
7.7628 | -91.0263
7.7631 | -91.0265
7.7635 | -91.0268
7.7638 | -91.0274
7.7642 | -91.0277
7.7645 | -91.0282
7.7648 | -91.0288
7.7652 | -91.0308
7.7655 | -91.0328
7.7659 | -91.0333
7.7662 | -91.0339
7.7665 | -91.0347
7.7669 | -91.0351
7.7672 | -91.0358
7.7676 | -91.0363
7.7679 | -91.0367
7.7682 | -91.0365
7.7686 | -91.0355
7.7689 | -91.035
7.7693 | -91.0352
7.7696 | -91.0359
7.7699 | -91.0369
7.7703 | -91.0377
7.7706 | -91.0383
7.771 | -91.0386
7.7713 | -91.039
7.7716 | -91.0393
7.772 | -91.0402
7.7723 | -91.0411
7.7727 | -91.0417
7.773 | -91.0421
7.7733 | -91.0427
7.7737 | -91.0432
7.774 | -91.0437
7.7744 | -91.0443
7.7747 | -91.0453
7.775 | -91.0458
7.7754 | -91.0466
7.7757 | -91.0481
7.7761 | -91.0488
7.7764 | -91.0496
7.7767 | -91.0507
7.7771 | -91.0526
7.7774 | -91.0537
7.7778 | -91.0545
7.7781 | -91.0551
7.7785 | -91.0558
7.7788 | -91.0556
7.7791 | -91.0563
7.7795 | -91.0571
7.7798 | -91.0575
7.7802 | -91.0581
7.7805 | -91.0589
7.7808 | -91.0592
7.7812 | -91.0596
7.7815 | -91.061
7.7819 | -91.0622
7.7822 | -91.0627
7.7825 | -91.0632
7.7829 | -91.0635
7.7832 | -91.0642
7.7836 | -91.0647
7.7839 | -91.0651
7.7842 | -91.0663
7.7846 | -91.0679
7.7849 | -91.069
7.7853 | -91.0698
7.7856 | -91.0703
7.7859 | -91.0709
7.7863 | -91.0714
7.7866 | -91.0718
7.787 | -91.0735
7.7873 | -91.0751
7.7876 | -91.0757
7.788 | -91.0756
7.7883 | -91.0758
7.7887 | -91.0762
7.789 | -91.0767
7.7893 | -91.0771
7.7897 | -91.0775
7.79 | -91.0776
7.7904 | -91.0781
7.7907 | -91.0785
7.791 | -91.0796
7.7914 | -91.0806
7.7917 | -91.0812
7.7921 | -91.0817
7.7924 | -91.0818
7.7928 | -91.0821
7.7931 | -91.0826
7.7934 | -91.0833
7.7938 | -91.0833
7.7941 | -91.0832
7.7945 | -91.0844
7.7948 | -91.0847
7.7951 | -91.0847
7.7955 | -91.0848
7.7958 | -91.0859
7.7962 | -91.0866
7.7965 | -91.0874
7.7968 | -91.0882
7.7972 | -91.0903
7.7975 | -91.0921
7.7979 | -91.0934
7.7982 | -91.0942
7.7985 | -91.0947
7.7989 | -91.0951
7.7992 | -91.0961
7.7996 | -91.0964
7.7999 | -91.0972
7.8002 | -91.0972
7.8006 | -91.0977
7.8009 | -91.0983
7.8013 | -91.0986
7.8016 | -91.0992
7.8019 | -91.1003
7.8023 | -91.1012
7.8026 | -91.1014
7.803 | -91.102
7.8033 | -91.1024
7.8036 | -91.1028
7.804 | -91.1036
7.8043 | -91.1049
7.8047 | -91.1068
7.805 | -91.1081
7.8053 | -91.1095
7.8057 | -91.1107
7.806 | -91.1113
7.8064 | -91.1122
7.8067 | -91.114
7.8071 | -91.1153
7.8074 | -91.1157
7.8077 | -91.1173
7.8081 | -91.1186
7.8084 | -91.1192
7.8088 | -91.1208
7.8091 | -91.1229
7.8094 | -91.1235
7.8098 | -91.124
7.8101 | -91.1242
7.8105 | -91.1234
7.8108 | -91.1221
7.8111 | -91.1217
7.8115 | -91.1224
7.8118 | -91.1231
7.8122 | -91.1234
7.8125 | -91.1237
7.8128 | -91.1241
7.8132 | -91.1247
7.8135 | -91.1252
7.8139 | -91.1253
7.8142 | -91.1258
7.8145 | -91.1265
7.8149 | -91.1272
7.8152 | -91.1279
7.8156 | -91.1287
7.8159 | -91.1292
7.8162 | -91.1296
7.8166 | -91.1299
7.8169 | -91.13
7.8173 | -91.1302
7.8176 | -91.1307
7.8179 | -91.1306
7.8183 | -91.1304
7.8186 | -91.131
7.819 | -91.1316
7.8193 | -91.1325
7.8197 | -91.1332
7.82 | -91.1343
7.8203 | -91.1342
7.8207 | -91.1345
7.821 | -91.135
7.8214 | -91.1358
7.8217 | -91.136
7.822 | -91.1364
7.8224 | -91.1371
7.8227 | -91.1366
7.8231 | -91.1368
7.8234 | -91.1376
7.8237 | -91.138
7.8241 | -91.1384
7.8244 | -91.1387
7.8248 | -91.1386
7.8251 | -91.1382
7.8254 | -91.138
7.8258 | -91.1385
7.8261 | -91.139
7.8265 | -91.1379
7.8268 | -91.1377
7.8271 | -91.1381
7.8275 | -91.1384
7.8278 | -91.1394
7.8282 | -91.141
7.8285 | -91.1416
7.8288 | -91.1424
7.8292 | -91.1433
7.8295 | -91.1442
7.8299 | -91.1448
7.8302 | -91.1451
7.8305 | -91.1457
7.8309 | -91.1463
7.8312 | -91.1466
7.8316 | -91.1465
7.8319 | -91.1462
7.8322 | -91.1467
7.8326 | -91.1474
7.8329 | -91.1486
7.8333 | -91.1498
7.8336 | -91.1504
7.834 | -91.1511
7.8343 | -91.152
7.8346 | -91.1526
7.835 | -91.1532
7.8353 | -91.1536
7.8357 | -91.1543
7.836 | -91.1554
7.8363 | -91.157
7.8367 | -91.1579
7.837 | -91.1583
7.8374 | -91.1588
7.8377 | -91.1602
7.838 | -91.1617
7.8384 | -91.1625
7.8387 | -91.163
7.8391 | -91.1637
7.8394 | -91.1646
7.8397 | -91.1648
7.8401 | -91.1653
7.8404 | -91.1658
7.8408 | -91.1664
7.8411 | -91.1671
7.8414 | -91.1676
7.8418 | -91.1681
7.8421 | -91.1689
7.8425 | -91.1692
7.8428 | -91.1701
7.8431 | -91.1715
7.8435 | -91.173
7.8438 | -91.1738
7.8442 | -91.1742
7.8445 | -91.175
7.8448 | -91.1759
7.8452 | -91.1777
7.8455 | -91.1808
7.8459 | -91.1826
7.8462 | -91.1828
7.8465 | -91.1836
7.8469 | -91.1841
7.8472 | -91.1848
7.8476 | -91.1845
7.8479 | -91.1845
7.8483 | -91.1854
7.8486 | -91.1863
7.8489 | -91.1869
7.8493 | -91.1871
7.8496 | -91.1874
7.85 | -91.1888
7.8503 | -91.1905
7.8506 | -91.1903
7.851 | -91.1906
7.8513 | -91.1914
7.8517 | -91.1916
7.852 | -91.1921
7.8523 | -91.1926
7.8527 | -91.193
7.853 | -91.1933
7.8534 | -91.1939
7.8537 | -91.1943
7.854 | -91.1945
7.8544 | -91.1954
7.8547 | -91.1963
7.8551 | -91.1965
7.8554 | -91.1971
7.8557 | -91.1974
7.8561 | -91.1979
7.8564 | -91.1988
7.8568 | -91.1995
7.8571 | -91.2001
7.8574 | -91.2006
7.8578 | -91.2006
7.8581 | -91.2002
7.8585 | -91.2003
7.8588 | -91.201
7.8591 | -91.2014
7.8595 | -91.2019
7.8598 | -91.2025
7.8602 | -91.2031
7.8605 | -91.2037
7.8608 | -91.2041
7.8612 | -91.2038
7.8615 | -91.2041
7.8619 | -91.2044
7.8622 | -91.205
7.8626 | -91.2061
7.8629 | -91.2074
7.8632 | -91.2078
7.8636 | -91.2083
7.8639 | -91.2088
7.8643 | -91.2093
7.8646 | -91.2099
7.8649 | -91.2105
7.8653 | -91.212
7.8656 | -91.2139
7.866 | -91.2149
7.8663 | -91.2155
7.8666 | -91.2161
7.867 | -91.2168
7.8673 | -91.2176
7.8677 | -91.2182
7.868 | -91.2188
7.8683 | -91.2196
7.8687 | -91.2205
7.869 | -91.222
7.8694 | -91.2238
7.8697 | -91.225
7.87 | -91.226
7.8704 | -91.2273
7.8707 | -91.2283
7.8711 | -91.2289
7.8714 | -91.2298
7.8717 | -91.2306
7.8721 | -91.2317
7.8724 | -91.2327
7.8728 | -91.2329
7.8731 | -91.2334
7.8734 | -91.2343
7.8738 | -91.2351
7.8741 | -91.2362
7.8745 | -91.2381
7.8748 | -91.2395
7.8751 | -91.2398
7.8755 | -91.2411
7.8758 | -91.2417
7.8762 | -91.2418
7.8765 | -91.2424
7.8769 | -91.2431
7.8772 | -91.2439
7.8775 | -91.2446
7.8779 | -91.2449
7.8782 | -91.2462
7.8786 | -91.2464
7.8789 | -91.2465
7.8792 | -91.2472
7.8796 | -91.2484
7.8799 | -91.2494
7.8803 | -91.2498
7.8806 | -91.2503
7.8809 | -91.2508
7.8813 | -91.2513
7.8816 | -91.252
7.882 | -91.2526
7.8823 | -91.2532
7.8826 | -91.2538
7.883 | -91.2545
7.8833 | -91.2557
7.8837 | -91.2573
7.884 | -91.2579
7.8843 | -91.2585
7.8847 | -91.2592
7.885 | -91.2593
7.8854 | -91.2593
7.8857 | -91.2595
7.886 | -91.2608
7.8864 | -91.2616
7.8867 | -91.262
7.8871 | -91.2624
7.8874 | -91.2632
7.8877 | -91.2636
7.8881 | -91.2642
7.8884 | -91.2651
7.8888 | -91.2658
7.8891 | -91.2668
7.8895 | -91.2673
7.8898 | -91.2679
7.8901 | -91.2683
7.8905 | -91.2685
7.8908 | -91.2693
7.8912 | -91.2701
7.8915 | -91.2713
7.8918 | -91.2722
7.8922 | -91.2723
7.8925 | -91.2728
7.8929 | -91.2737
7.8932 | -91.2741
7.8935 | -91.275
7.8939 | -91.2766
7.8942 | -91.2765
7.8946 | -91.2763
7.8949 | -91.2773
7.8952 | -91.2787
7.8956 | -91.2793
7.8959 | -91.2801
7.8963 | -91.2808
7.8966 | -91.2818
7.8969 | -91.2831
7.8973 | -91.2841
7.8976 | -91.2845
7.898 | -91.2851
7.8983 | -91.2851
7.8986 | -91.2858
7.899 | -91.2865
7.8993 | -91.2871
7.8997 | -91.2872
7.9 | -91.2869
7.9003 | -91.2879
7.9007 | -91.289
7.901 | -91.2897
7.9014 | -91.2902
7.9017 | -91.2909
7.902 | -91.2916
7.9024 | -91.2923
7.9027 | -91.2931
7.9031 | -91.2942
7.9034 | -91.2948
7.9038 | -91.295
7.9041 | -91.2949
7.9044 | -91.2954
7.9048 | -91.2962
7.9051 | -91.2969
7.9055 | -91.2971
7.9058 | -91.2983
7.9061 | -91.2994
7.9065 | -91.2999
7.9068 | -91.3005
7.9072 | -91.3012
7.9075 | -91.3021
7.9078 | -91.303
7.9082 | -91.3038
7.9085 | -91.305
7.9089 | -91.3054
7.9092 | -91.3061
7.9095 | -91.3072
7.9099 | -91.3081
7.9102 | -91.3088
7.9106 | -91.3093
7.9109 | -91.3099
7.9112 | -91.311
7.9116 | -91.312
7.9119 | -91.3124
7.9123 | -91.313
7.9126 | -91.3135
7.9129 | -91.3147
7.9133 | -91.3157
7.9136 | -91.3166
7.914 | -91.3173
7.9143 | -91.318
7.9146 | -91.3192
7.915 | -91.3198
7.9153 | -91.3204
7.9157 | -91.321
7.916 | -91.3215
7.9163 | -91.3219
7.9167 | -91.3225
7.917 | -91.3233
7.9174 | -91.3239
7.9177 | -91.3246
7.9181 | -91.3255
7.9184 | -91.3257
7.9187 | -91.3265
7.9191 | -91.3275
7.9194 | -91.3282
7.9198 | -91.3298
7.9201 | -91.3307
7.9204 | -91.3311
7.9208 | -91.3316
7.9211 | -91.3321
7.9215 | -91.333
7.9218 | -91.3331
7.9221 | -91.3336
7.9225 | -91.3343
7.9228 | -91.3346
7.9232 | -91.3355
7.9235 | -91.3362
7.9238 | -91.3368
7.9242 | -91.3374
7.9245 | -91.338
7.9249 | -91.3383
7.9252 | -91.3386
7.9255 | -91.3391
7.9259 | -91.3401
7.9262 | -91.341
7.9266 | -91.3419
7.9269 | -91.3425
7.9272 | -91.3431
7.9276 | -91.3427
7.9279 | -91.3431
7.9283 | -91.3453
7.9286 | -91.3469
7.9289 | -91.3476
7.9293 | -91.3483
7.9296 | -91.3486
7.93 | -91.3491
7.9303 | -91.3499
7.9306 | -91.3506
7.931 | -91.3511
7.9313 | -91.3516
7.9317 | -91.3522
7.932 | -91.3528
7.9324 | -91.3535
7.9327 | -91.3539
7.933 | -91.3541
7.9334 | -91.3536
7.9337 | -91.3528
7.9341 | -91.3528
7.9344 | -91.353
7.9347 | -91.354
7.9351 | -91.3554
7.9354 | -91.357
7.9358 | -91.3586
7.9361 | -91.359
7.9364 | -91.3591
7.9368 | -91.3592
7.9371 | -91.3604
7.9375 | -91.3613
7.9378 | -91.3627
7.9381 | -91.3633
7.9385 | -91.364
7.9388 | -91.3646
7.9392 | -91.3653
7.9395 | -91.3663
7.9398 | -91.3668
7.9402 | -91.3674
7.9405 | -91.3682
7.9409 | -91.3694
7.9412 | -91.3707
7.9415 | -91.3714
7.9419 | -91.3721
7.9422 | -91.3725
7.9426 | -91.3732
7.9429 | -91.3735
7.9432 | -91.3743
7.9436 | -91.3756
7.9439 | -91.3762
7.9443 | -91.3775
7.9446 | -91.3789
7.945 | -91.38
7.9453 | -91.3805
7.9456 | -91.3811
7.946 | -91.3818
7.9463 | -91.3824
7.9467 | -91.3828
7.947 | -91.3835
7.9473 | -91.3842
7.9477 | -91.3848
7.948 | -91.3855
7.9484 | -91.3866
7.9487 | -91.388
7.949 | -91.3885
7.9494 | -91.3897
7.9497 | -91.3905
7.9501 | -91.3913
7.9504 | -91.3921
7.9507 | -91.3929
7.9511 | -91.3937
7.9514 | -91.3946
7.9518 | -91.3957
7.9521 | -91.397
7.9524 | -91.3978
7.9528 | -91.3999
7.9531 | -91.4021
7.9535 | -91.4031
7.9538 | -91.4038
7.9541 | -91.4049
7.9545 | -91.406
7.9548 | -91.4064
7.9552 | -91.4071
7.9555 | -91.4077
7.9558 | -91.4083
7.9562 | -91.4089
7.9565 | -91.4102
7.9569 | -91.4105
7.9572 | -91.4109
7.9575 | -91.4111
7.9579 | -91.412
7.9582 | -91.4126
7.9586 | -91.4133
7.9589 | -91.4143
7.9592 | -91.4156
7.9596 | -91.4163
7.9599 | -91.4167
7.9603 | -91.4171
7.9606 | -91.4175
7.961 | -91.4179
7.9613 | -91.4188
7.9616 | -91.4195
7.962 | -91.4202
7.9623 | -91.4214
7.9627 | -91.4222
7.963 | -91.4227
7.9633 | -91.4244
7.9637 | -91.4257
7.964 | -91.4262
7.9644 | -91.4266
7.9647 | -91.4276
7.965 | -91.4286
7.9654 | -91.4298
7.9657 | -91.4303
7.9661 | -91.4317
7.9664 | -91.4324
7.9667 | -91.4327
7.9671 | -91.433
7.9674 | -91.4337
7.9678 | -91.4342
7.9681 | -91.4352
7.9684 | -91.4359
7.9688 | -91.4368
7.9691 | -91.437
7.9695 | -91.4365
7.9698 | -91.437
7.9701 | -91.4376
7.9705 | -91.4382
7.9708 | -91.4391
7.9712 | -91.4398
7.9715 | -91.4404
7.9718 | -91.4413
7.9722 | -91.4418
7.9725 | -91.4421
7.9729 | -91.4428
7.9732 | -91.4437
7.9736 | -91.445
7.9739 | -91.4462
7.9742 | -91.4469
7.9746 | -91.4474
7.9749 | -91.4478
7.9753 | -91.4488
7.9756 | -91.4496
7.9759 | -91.45
7.9763 | -91.4505
7.9766 | -91.4512
7.977 | -91.4523
7.9773 | -91.4534
7.9776 | -91.4544
7.978 | -91.4545
7.9783 | -91.4551
7.9787 | -91.4559
7.979 | -91.4574
7.9793 | -91.4596
7.9797 | -91.4606
7.98 | -91.4608
7.9804 | -91.4616
7.9807 | -91.4626
7.981 | -91.4632
7.9814 | -91.4637
7.9817 | -91.4648
7.9821 | -91.4659
7.9824 | -91.4666
7.9827 | -91.4672
7.9831 | -91.4676
7.9834 | -91.4679
7.9838 | -91.4684
7.9841 | -91.4697
7.9844 | -91.4722
7.9848 | -91.4736
7.9851 | -91.4744
7.9855 | -91.4748
7.9858 | -91.4752
7.9861 | -91.4759
7.9865 | -91.4768
7.9868 | -91.4777
7.9872 | -91.4793
7.9875 | -91.4797
7.9879 | -91.4802
7.9882 | -91.4806
7.9885 | -91.4813
7.9889 | -91.4819
7.9892 | -91.4825
7.9896 | -91.4833
7.9899 | -91.4838
7.9902 | -91.4849
7.9906 | -91.4858
7.9909 | -91.4862
7.9913 | -91.4867
7.9916 | -91.4874
7.9919 | -91.488
7.9923 | -91.4885
7.9926 | -91.4891
7.993 | -91.4896
7.9933 | -91.4897
7.9936 | -91.4895
7.994 | -91.4899
7.9943 | -91.4905
7.9947 | -91.4913
7.995 | -91.4923
7.9953 | -91.4928
7.9957 | -91.4935
7.996 | -91.4948
7.9964 | -91.4955
7.9967 | -91.4963
7.997 | -91.497
7.9974 | -91.4978
7.9977 | -91.4988
7.9981 | -91.4992
7.9984 | -91.5
7.9987 | -91.5015
7.9991 | -91.5021
7.9994 | -91.5023
7.9998 | -91.5028
8.0001 | -91.5034
8.0004 | -91.5041
8.0008 | -91.5049
8.0011 | -91.5063
8.0015 | -91.507
8.0018 | -91.5074
8.0022 | -91.508
8.0025 | -91.5087
8.0028 | -91.5092
8.0032 | -91.5101
8.0035 | -91.5114
8.0039 | -91.5122
8.0042 | -91.5129
8.0045 | -91.5134
8.0049 | -91.5143
8.0052 | -91.516
8.0056 | -91.5173
8.0059 | -91.5182
8.0062 | -91.5195
8.0066 | -91.5204
8.0069 | -91.5212
8.0073 | -91.5218
8.0076 | -91.5228
8.0079 | -91.5242
8.0083 | -91.5253
8.0086 | -91.5265
8.009 | -91.5275
8.0093 | -91.5282
8.0096 | -91.5289
8.01 | -91.5295
8.0103 | -91.5303
8.0107 | -91.531
8.011 | -91.532
8.0113 | -91.5331
8.0117 | -91.5336
8.012 | -91.5341
8.0124 | -91.5349
8.0127 | -91.536
8.013 | -91.5369
8.0134 | -91.5375
8.0137 | -91.5384
8.0141 | -91.5389
8.0144 | -91.5391
8.0147 | -91.5395
8.0151 | -91.5401
8.0154 | -91.541
8.0158 | -91.5414
8.0161 | -91.5426
8.0165 | -91.5434
8.0168 | -91.5439
8.0171 | -91.5446
8.0175 | -91.5454
8.0178 | -91.546
8.0182 | -91.5468
8.0185 | -91.5475
8.0188 | -91.5482
8.0192 | -91.5489
8.0195 | -91.5494
8.0199 | -91.5496
8.0202 | -91.5503
8.0205 | -91.5512
8.0209 | -91.5517
8.0212 | -91.5515
8.0216 | -91.5511
8.0219 | -91.5509
8.0222 | -91.5516
8.0226 | -91.5523
8.0229 | -91.5524
8.0233 | -91.5532
8.0236 | -91.5539
8.0239 | -91.5543
8.0243 | -91.5546
8.0246 | -91.5558
8.025 | -91.557
8.0253 | -91.5579
8.0256 | -91.5586
8.026 | -91.5595
8.0263 | -91.5602
8.0267 | -91.5612
8.027 | -91.5617
8.0273 | -91.5612
8.0277 | -91.5607
8.028 | -91.5612
8.0284 | -91.5621
8.0287 | -91.5636
8.029 | -91.5642
8.0294 | -91.565
8.0297 | -91.5665
8.0301 | -91.5682
8.0304 | -91.5689
8.0308 | -91.5691
8.0311 | -91.5696
8.0314 | -91.5696
8.0318 | -91.5693
8.0321 | -91.5697
8.0325 | -91.5705
8.0328 | -91.5708
8.0331 | -91.5713
8.0335 | -91.5719
8.0338 | -91.5728
8.0342 | -91.5737
8.0345 | -91.5751
8.0348 | -91.5759
8.0352 | -91.5766
8.0355 | -91.5779
8.0359 | -91.5792
8.0362 | -91.5801
8.0365 | -91.5804
8.0369 | -91.581
8.0372 | -91.5823
8.0376 | -91.5834
8.0379 | -91.5841
8.0382 | -91.585
8.0386 | -91.5857
8.0389 | -91.5867
8.0393 | -91.5874
8.0396 | -91.5879
8.0399 | -91.5887
8.0403 | -91.5897
8.0406 | -91.5908
8.041 | -91.5915
8.0413 | -91.5921
8.0416 | -91.5929
8.042 | -91.5936
8.0423 | -91.5944
8.0427 | -91.5954
8.043 | -91.5962
8.0434 | -91.5971
8.0437 | -91.598
8.044 | -91.5982
8.0444 | -91.599
8.0447 | -91.5999
8.0451 | -91.6007
8.0454 | -91.6019
8.0457 | -91.6033
8.0461 | -91.6047
8.0464 | -91.6057
8.0468 | -91.6066
8.0471 | -91.6073
8.0474 | -91.608
8.0478 | -91.6091
8.0481 | -91.6117
8.0485 | -91.6139
8.0488 | -91.6153
8.0491 | -91.6159
8.0495 | -91.6164
8.0498 | -91.6174
8.0502 | -91.6181
8.0505 | -91.6192
8.0508 | -91.6203
8.0512 | -91.6203
8.0515 | -91.62
8.0519 | -91.6193
8.0522 | -91.6196
8.0525 | -91.6211
8.0529 | -91.6226
8.0532 | -91.6237
8.0536 | -91.6247
8.0539 | -91.6256
8.0542 | -91.6264
8.0546 | -91.6275
8.0549 | -91.6284
8.0553 | -91.629
8.0556 | -91.63
8.0559 | -91.6299
8.0563 | -91.63
8.0566 | -91.6308
8.057 | -91.632
8.0573 | -91.6326
8.0577 | -91.6329
8.058 | -91.6336
8.0583 | -91.6351
8.0587 | -91.6363
8.059 | -91.6371
8.0594 | -91.6376
8.0597 | -91.6382
8.06 | -91.6393
8.0604 | -91.6398
8.0607 | -91.6403
8.0611 | -91.6411
8.0614 | -91.642
8.0617 | -91.643
8.0621 | -91.6438
8.0624 | -91.6446
8.0628 | -91.6465
8.0631 | -91.6483
8.0634 | -91.6493
8.0638 | -91.65
8.0641 | -91.652
8.0645 | -91.653
8.0648 | -91.6537
8.0651 | -91.6548
8.0655 | -91.6549
8.0658 | -91.6558
8.0662 | -91.6579
8.0665 | -91.6591
8.0668 | -91.6598
8.0672 | -91.6606
8.0675 | -91.6613
8.0679 | -91.6614
8.0682 | -91.6611
8.0685 | -91.6606
8.0689 | -91.6608
8.0692 | -91.6613
8.0696 | -91.6612
8.0699 | -91.6616
8.0702 | -91.6628
8.0706 | -91.6629
8.0709 | -91.6638
8.0713 | -91.6655
8.0716 | -91.6664
8.072 | -91.668
8.0723 | -91.6702
8.0726 | -91.6719
8.073 | -91.6728
8.0733 | -91.6737
8.0737 | -91.6748
8.074 | -91.6758
8.0743 | -91.6763
8.0747 | -91.677
8.075 | -91.6782
8.0754 | -91.6785
8.0757 | -91.6795
8.076 | -91.6798
8.0764 | -91.6809
8.0767 | -91.6818
8.0771 | -91.6825
8.0774 | -91.6831
8.0777 | -91.6848
8.0781 | -91.6864
8.0784 | -91.6871
8.0788 | -91.6878
8.0791 | -91.6895
8.0794 | -91.6906
8.0798 | -91.6918
8.0801 | -91.6925
8.0805 | -91.693
8.0808 | -91.6949
8.0811 | -91.6959
8.0815 | -91.6962
8.0818 | -91.6969
8.0822 | -91.6976
8.0825 | -91.6979
8.0828 | -91.6983
8.0832 | -91.6989
8.0835 | -91.6997
8.0839 | -91.7004
8.0842 | -91.7011
8.0846 | -91.7019
8.0849 | -91.7032
8.0852 | -91.7039
8.0856 | -91.7048
8.0859 | -91.706
8.0863 | -91.7071
8.0866 | -91.7079
8.0869 | -91.7086
8.0873 | -91.7092
8.0876 | -91.7097
8.088 | -91.7104
8.0883 | -91.7109
8.0886 | -91.7115
8.089 | -91.7122
8.0893 | -91.7126
8.0897 | -91.713
8.09 | -91.7137
8.0903 | -91.7148
8.0907 | -91.7157
8.091 | -91.7167
8.0914 | -91.7175
8.0917 | -91.7184
8.092 | -91.7191
8.0924 | -91.7202
8.0927 | -91.7217
8.0931 | -91.7235
8.0934 | -91.7252
8.0937 | -91.7261
8.0941 | -91.7264
8.0944 | -91.727
8.0948 | -91.7281
8.0951 | -91.7289
8.0954 | -91.7298
8.0958 | -91.7307
8.0961 | -91.7314
8.0965 | -91.7321
8.0968 | -91.7331
8.0971 | -91.7336
8.0975 | -91.7346
8.0978 | -91.7364
8.0982 | -91.7376
8.0985 | -91.7384
8.0989 | -91.7393
8.0992 | -91.7403
8.0995 | -91.7406
8.0999 | -91.7412
8.1002 | -91.742
8.1006 | -91.7429
8.1009 | -91.7439
8.1012 | -91.7447
8.1016 | -91.7454
8.1019 | -91.7461
8.1023 | -91.7467
8.1026 | -91.7474
8.1029 | -91.7485
8.1033 | -91.7495
8.1036 | -91.7503
8.104 | -91.7516
8.1043 | -91.7524
8.1046 | -91.753
8.105 | -91.7539
8.1053 | -91.7549
8.1057 | -91.7552
8.106 | -91.756
8.1063 | -91.7565
8.1067 | -91.7574
8.107 | -91.7582
8.1074 | -91.7583
8.1077 | -91.7588
8.108 | -91.7599
8.1084 | -91.762
8.1087 | -91.7631
8.1091 | -91.7634
8.1094 | -91.7638
8.1097 | -91.7642
8.1101 | -91.7646
8.1104 | -91.765
8.1108 | -91.7657
8.1111 | -91.7664
8.1114 | -91.767
8.1118 | -91.7677
8.1121 | -91.7692
8.1125 | -91.7702
8.1128 | -91.7708
8.1132 | -91.7716
8.1135 | -91.7727
8.1138 | -91.7735
8.1142 | -91.7744
8.1145 | -91.7757
8.1149 | -91.7762
8.1152 | -91.777
8.1155 | -91.779
8.1159 | -91.7804
8.1162 | -91.7821
8.1166 | -91.7835
8.1169 | -91.7843
8.1172 | -91.7849
8.1176 | -91.7857
8.1179 | -91.7864
8.1183 | -91.7882
8.1186 | -91.7904
8.1189 | -91.7915
8.1193 | -91.792
8.1196 | -91.7925
8.12 | -91.7941
8.1203 | -91.7954
8.1206 | -91.7959
8.121 | -91.7969
8.1213 | -91.798
8.1217 | -91.7988
8.122 | -91.7998
8.1223 | -91.8007
8.1227 | -91.8016
8.123 | -91.8023
8.1234 | -91.8037
8.1237 | -91.8046
8.124 | -91.8057
8.1244 | -91.8063
8.1247 | -91.8073
8.1251 | -91.8084
8.1254 | -91.809
8.1257 | -91.8095
8.1261 | -91.8104
8.1264 | -91.8111
8.1268 | -91.8116
8.1271 | -91.8107
8.1275 | -91.8115
8.1278 | -91.8125
8.1281 | -91.8134
8.1285 | -91.8135
8.1288 | -91.8147
8.1292 | -91.8154
8.1295 | -91.8162
8.1298 | -91.8173
8.1302 | -91.8181
8.1305 | -91.8187
8.1309 | -91.8193
8.1312 | -91.8198
8.1315 | -91.8205
8.1319 | -91.8199
8.1322 | -91.8198
8.1326 | -91.8198
8.1329 | -91.8208
8.1332 | -91.8229
8.1336 | -91.8248
8.1339 | -91.8247
8.1343 | -91.8258
8.1346 | -91.8254
8.1349 | -91.8272
8.1353 | -91.8281
8.1356 | -91.8286
8.136 | -91.828
8.1363 | -91.8273
8.1366 | -91.8286
8.137 | -91.8301
8.1373 | -91.8302
8.1377 | -91.8286
8.138 | -91.8279
8.1383 | -91.8278
8.1387 | -91.827
8.139 | -91.827
8.1394 | -91.8267
8.1397 | -91.8256
8.1401 | -91.8251
8.1404 | -91.8242
8.1407 | -91.8243
8.1411 | -91.8235
8.1414 | -91.8239
8.1418 | -91.8235
8.1421 | -91.8244
8.1424 | -91.8249
8.1428 | -91.8256
8.1431 | -91.8267
8.1435 | -91.8279
8.1438 | -91.8285
8.1441 | -91.8289
8.1445 | -91.8297
8.1448 | -91.8306
8.1452 | -91.8316
8.1455 | -91.8322
8.1458 | -91.8337
8.1462 | -91.8353
8.1465 | -91.8365
8.1469 | -91.8374
8.1472 | -91.8387
8.1475 | -91.8396
8.1479 | -91.8409
8.1482 | -91.842
8.1486 | -91.8425
8.1489 | -91.8433
8.1492 | -91.8438
8.1496 | -91.8449
8.1499 | -91.8459
8.1503 | -91.847
8.1506 | -91.8482
8.1509 | -91.8487
8.1513 | -91.849
8.1516 | -91.8491
8.152 | -91.8498
8.1523 | -91.851
8.1526 | -91.8527
8.153 | -91.8537
8.1533 | -91.8544
8.1537 | -91.8551
8.154 | -91.8553
8.1544 | -91.8556
8.1547 | -91.8563
8.155 | -91.8565
8.1554 | -91.8571
8.1557 | -91.858
8.1561 | -91.8592
8.1564 | -91.8604
8.1567 | -91.8611
8.1571 | -91.8615
8.1574 | -91.8623
8.1578 | -91.8638
8.1581 | -91.8642
8.1584 | -91.8647
8.1588 | -91.8658
8.1591 | -91.8664
8.1595 | -91.867
8.1598 | -91.8675
8.1601 | -91.8681
8.1605 | -91.8691
8.1608 | -91.8706
8.1612 | -91.8716
8.1615 | -91.8724
8.1618 | -91.873
8.1622 | -91.8738
8.1625 | -91.8753
8.1629 | -91.8761
8.1632 | -91.8766
8.1635 | -91.8765
8.1639 | -91.8764
8.1642 | -91.8775
8.1646 | -91.878
8.1649 | -91.8787
8.1652 | -91.8795
8.1656 | -91.8805
8.1659 | -91.8815
8.1663 | -91.8823
8.1666 | -91.883
8.1669 | -91.8837
8.1673 | -91.8845
8.1676 | -91.8851
8.168 | -91.8857
8.1683 | -91.8862
8.1687 | -91.8867
8.169 | -91.888
8.1693 | -91.8889
8.1697 | -91.8894
8.17 | -91.8899
8.1704 | -91.8901
8.1707 | -91.8902
8.171 | -91.8903
8.1714 | -91.8897
8.1717 | -91.889
8.1721 | -91.8894
8.1724 | -91.89
8.1727 | -91.8903
8.1731 | -91.8905
8.1734 | -91.89
8.1738 | -91.8907
8.1741 | -91.8913
8.1744 | -91.8916
8.1748 | -91.892
8.1751 | -91.8924
8.1755 | -91.8928
8.1758 | -91.8934
8.1761 | -91.8942
8.1765 | -91.8946
8.1768 | -91.895
8.1772 | -91.8959
8.1775 | -91.8965
8.1778 | -91.8973
8.1782 | -91.8983
8.1785 | -91.8989
8.1789 | -91.8996
8.1792 | -91.9002
8.1795 | -91.9007
8.1799 | -91.9013
8.1802 | -91.9015
8.1806 | -91.902
8.1809 | -91.9023
8.1812 | -91.9034
8.1816 | -91.9037
8.1819 | -91.9041
8.1823 | -91.9046
8.1826 | -91.9049
8.183 | -91.9046
8.1833 | -91.9041
8.1836 | -91.9047
8.184 | -91.905
8.1843 | -91.9056
8.1847 | -91.9064
8.185 | -91.9069
8.1853 | -91.9073
8.1857 | -91.9084
8.186 | -91.9106
8.1864 | -91.9126
8.1867 | -91.9133
8.187 | -91.914
8.1874 | -91.9145
8.1877 | -91.9148
8.1881 | -91.9156
8.1884 | -91.9166
8.1887 | -91.9175
8.1891 | -91.9175
8.1894 | -91.9167
8.1898 | -91.9167
8.1901 | -91.9168
8.1904 | -91.9172
8.1908 | -91.9174
8.1911 | -91.9176
8.1915 | -91.9179
8.1918 | -91.9187
8.1921 | -91.9196
8.1925 | -91.9207
8.1928 | -91.9208
8.1932 | -91.9214
8.1935 | -91.9218
8.1938 | -91.9221
8.1942 | -91.9223
8.1945 | -91.9221
8.1949 | -91.922
8.1952 | -91.9221
8.1955 | -91.9229
8.1959 | -91.9237
8.1962 | -91.924
8.1966 | -91.9243
8.1969 | -91.9245
8.1973 | -91.9245
8.1976 | -91.9248
8.1979 | -91.9251
8.1983 | -91.9252
8.1986 | -91.9255
8.199 | -91.9265
8.1993 | -91.9271
8.1996 | -91.9267
8.2 | -91.9268
8.2003 | -91.9271
8.2007 | -91.9273
8.201 | -91.927
8.2013 | -91.9271
8.2017 | -91.9278
8.202 | -91.9282
8.2024 | -91.9286
8.2027 | -91.929
8.203 | -91.9293
8.2034 | -91.9298
8.2037 | -91.9303
8.2041 | -91.9313
8.2044 | -91.9318
8.2047 | -91.9322
8.2051 | -91.9341
8.2054 | -91.9361
8.2058 | -91.9366
8.2061 | -91.9369
8.2064 | -91.9372
8.2068 | -91.9376
8.2071 | -91.9381
8.2075 | -91.9388
8.2078 | -91.9391
8.2081 | -91.9391
8.2085 | -91.9391
8.2088 | -91.9395
8.2092 | -91.9399
8.2095 | -91.9406
8.2098 | -91.9408
8.2102 | -91.9411
8.2105 | -91.9413
8.2109 | -91.9419
8.2112 | -91.9421
8.2116 | -91.9412
8.2119 | -91.9411
8.2122 | -91.9413
8.2126 | -91.9412
8.2129 | -91.9415
8.2133 | -91.9414
8.2136 | -91.9412
8.2139 | -91.9415
8.2143 | -91.9418
8.2146 | -91.9424
8.215 | -91.9427
8.2153 | -91.9433
8.2156 | -91.9431
8.216 | -91.9433
8.2163 | -91.9441
8.2167 | -91.9441
8.217 | -91.9444
8.2173 | -91.9449
8.2177 | -91.9449
8.218 | -91.9447
8.2184 | -91.9451
8.2187 | -91.9455
8.219 | -91.9457
8.2194 | -91.9459
8.2197 | -91.9463
8.2201 | -91.9473
8.2204 | -91.9484
8.2207 | -91.9488
8.2211 | -91.9487
8.2214 | -91.9487
8.2218 | -91.9488
8.2221 | -91.9488
8.2224 | -91.9489
8.2228 | -91.9491
8.2231 | -91.9492
8.2235 | -91.9496
8.2238 | -91.9502
8.2241 | -91.9509
8.2245 | -91.9514
8.2248 | -91.9514
8.2252 | -91.9517
8.2255 | -91.9508
8.2259 | -91.9506
8.2262 | -91.9512
8.2265 | -91.9509
8.2269 | -91.9502
8.2272 | -91.9506
8.2276 | -91.9508
8.2279 | -91.9514
8.2282 | -91.9524
8.2286 | -91.9535
8.2289 | -91.9536
8.2293 | -91.954
8.2296 | -91.9539
8.2299 | -91.954
8.2303 | -91.9537
8.2306 | -91.9537
8.231 | -91.9538
8.2313 | -91.9538
8.2316 | -91.9543
8.232 | -91.9547
8.2323 | -91.955
8.2327 | -91.9556
8.233 | -91.9556
8.2333 | -91.9554
8.2337 | -91.9558
8.234 | -91.9556
8.2344 | -91.9555
8.2347 | -91.9555
8.235 | -91.9553
8.2354 | -91.9554
8.2357 | -91.9555
8.2361 | -91.9556
8.2364 | -91.9559
8.2367 | -91.9559
8.2371 | -91.9556
8.2374 | -91.9557
8.2378 | -91.956
8.2381 | -91.956
8.2385 | -91.9561
8.2388 | -91.9559
8.2391 | -91.9557
8.2395 | -91.9553
8.2398 | -91.9554
8.2402 | -91.9564
8.2405 | -91.9569
8.2408 | -91.9578
8.2412 | -91.9577
8.2415 | -91.9576
8.2419 | -91.9579
8.2422 | -91.958
8.2425 | -91.9587
8.2429 | -91.9599
8.2432 | -91.9604
8.2436 | -91.96
8.2439 | -91.9597
8.2442 | -91.9596
8.2446 | -91.9596
8.2449 | -91.9595
8.2453 | -91.9595
8.2456 | -91.9593
8.2459 | -91.9594
8.2463 | -91.9607
8.2466 | -91.9622
8.247 | -91.9623
8.2473 | -91.9625
8.2476 | -91.9624
8.248 | -91.9622
8.2483 | -91.9619
8.2487 | -91.9618
8.249 | -91.9619
8.2493 | -91.9619
8.2497 | -91.9615
8.25 | -91.9615
8.2504 | -91.9613
8.2507 | -91.9608
8.251 | -91.9605
8.2514 | -91.9605
8.2517 | -91.9606
8.2521 | -91.9604
8.2524 | -91.9597
8.2528 | -91.9597
8.2531 | -91.96
8.2534 | -91.96
8.2538 | -91.9601
8.2541 | -91.9602
8.2545 | -91.96
8.2548 | -91.9599
8.2551 | -91.9601
8.2555 | -91.96
8.2558 | -91.96
8.2562 | -91.9596
8.2565 | -91.9594
8.2568 | -91.9594
8.2572 | -91.9601
8.2575 | -91.9608
8.2579 | -91.9614
8.2582 | -91.962
8.2585 | -91.9617
8.2589 | -91.9616
8.2592 | -91.9614
8.2596 | -91.9602
8.2599 | -91.9593
8.2602 | -91.959
8.2606 | -91.9591
8.2609 | -91.9597
8.2613 | -91.9602
8.2616 | -91.9603
8.2619 | -91.9605
8.2623 | -91.96
8.2626 | -91.9598
8.263 | -91.9597
8.2633 | -91.9602
8.2636 | -91.9599
8.264 | -91.959
8.2643 | -91.9595
8.2647 | -91.9596
8.265 | -91.9594
8.2653 | -91.9595
8.2657 | -91.9599
8.266 | -91.96
8.2664 | -91.9601
8.2667 | -91.9603
8.2671 | -91.9601
8.2674 | -91.96
8.2677 | -91.96
8.2681 | -91.9602
8.2684 | -91.9601
8.2688 | -91.96
8.2691 | -91.9602
8.2694 | -91.9607
8.2698 | -91.9613
8.2701 | -91.9612
8.2705 | -91.9611
8.2708 | -91.9612
8.2711 | -91.9616
8.2715 | -91.9603
8.2718 | -91.9599
8.2722 | -91.9598
8.2725 | -91.9582
8.2728 | -91.9575
8.2732 | -91.9565
8.2735 | -91.9557
8.2739 | -91.9552
8.2742 | -91.9545
8.2745 | -91.9545
8.2749 | -91.9542
8.2752 | -91.954
8.2756 | -91.9539
8.2759 | -91.9538
8.2762 | -91.9536
8.2766 | -91.9534
8.2769 | -91.9531
8.2773 | -91.9525
8.2776 | -91.9523
8.2779 | -91.9524
8.2783 | -91.9527
8.2786 | -91.9528
8.279 | -91.9533
8.2793 | -91.953
8.2797 | -91.9525
8.28 | -91.9529
8.2803 | -91.9533
8.2807 | -91.9532
8.281 | -91.953
8.2814 | -91.9529
8.2817 | -91.9523
8.282 | -91.9523
8.2824 | -91.9522
8.2827 | -91.9519
8.2831 | -91.9513
8.2834 | -91.9507
8.2837 | -91.951
8.2841 | -91.9506
8.2844 | -91.9511
8.2848 | -91.9512
8.2851 | -91.9511
8.2854 | -91.95
8.2858 | -91.9495
8.2861 | -91.9493
8.2865 | -91.9493
8.2868 | -91.9493
8.2871 | -91.9493
8.2875 | -91.9491
8.2878 | -91.9492
8.2882 | -91.9502
8.2885 | -91.9502
8.2888 | -91.9501
8.2892 | -91.9496
8.2895 | -91.9491
8.2899 | -91.9494
8.2902 | -91.9502
8.2905 | -91.9509
8.2909 | -91.9509
8.2912 | -91.9506
8.2916 | -91.9503
8.2919 | -91.95
8.2922 | -91.9497
8.2926 | -91.9491
8.2929 | -91.9489
8.2933 | -91.9484
8.2936 | -91.9477
8.294 | -91.9471
8.2943 | -91.9473
8.2946 | -91.9481
8.295 | -91.9481
8.2953 | -91.9478
8.2957 | -91.9477
8.296 | -91.9473
8.2963 | -91.9474
8.2967 | -91.9465
8.297 | -91.9465
8.2974 | -91.9462
8.2977 | -91.9463
8.298 | -91.9459
8.2984 | -91.9451
8.2987 | -91.9448
8.2991 | -91.9442
8.2994 | -91.9441
8.2997 | -91.9447
8.3001 | -91.9446
8.3004 | -91.9438
8.3008 | -91.9434
8.3011 | -91.9432
8.3014 | -91.9431
8.3018 | -91.9429
8.3021 | -91.9426
8.3025 | -91.9432
8.3028 | -91.9435
8.3031 | -91.9434
8.3035 | -91.9429
8.3038 | -91.9429
8.3042 | -91.9432
8.3045 | -91.9431
8.3048 | -91.9422
8.3052 | -91.9414
8.3055 | -91.9412
8.3059 | -91.9409
8.3062 | -91.9407
8.3065 | -91.9401
8.3069 | -91.9389
8.3072 | -91.9387
8.3076 | -91.9388
8.3079 | -91.9386
8.3083 | -91.9384
8.3086 | -91.9381
8.3089 | -91.9387
8.3093 | -91.9389
8.3096 | -91.938
8.31 | -91.9373
8.3103 | -91.937
8.3106 | -91.9367
8.311 | -91.9365
8.3113 | -91.9363
8.3117 | -91.9361
8.312 | -91.9356
8.3123 | -91.935
8.3127 | -91.9349
8.313 | -91.9341
8.3134 | -91.933
8.3137 | -91.9326
8.314 | -91.9324
8.3144 | -91.9324
8.3147 | -91.9316
8.3151 | -91.9308
8.3154 | -91.9305
8.3157 | -91.9308
8.3161 | -91.9313
8.3164 | -91.9314
8.3168 | -91.9318
8.3171 | -91.932
8.3174 | -91.9312
8.3178 | -91.9303
8.3181 | -91.9299
8.3185 | -91.9295
8.3188 | -91.9292
8.3191 | -91.9292
8.3195 | -91.9291
8.3198 | -91.9283
8.3202 | -91.9279
8.3205 | -91.9272
8.3208 | -91.9271
8.3212 | -91.9267
8.3215 | -91.9264
8.3219 | -91.9262
8.3222 | -91.926
8.3226 | -91.926
8.3229 | -91.9254
8.3232 | -91.9255
8.3236 | -91.9259
8.3239 | -91.926
8.3243 | -91.9261
8.3246 | -91.9257
8.3249 | -91.9253
8.3253 | -91.9258
8.3256 | -91.9257
8.326 | -91.9254
8.3263 | -91.9253
8.3266 | -91.926
8.327 | -91.9271
8.3273 | -91.9272
8.3277 | -91.9263
8.328 | -91.9259
8.3283 | -91.9254
8.3287 | -91.9254
8.329 | -91.9254
8.3294 | -91.9249
8.3297 | -91.9246
8.33 | -91.9242
8.3304 | -91.9244
8.3307 | -91.9241
8.3311 | -91.9236
8.3314 | -91.9231
8.3317 | -91.9204
8.3321 | -91.9186
8.3324 | -91.9183
8.3328 | -91.9174
8.3331 | -91.9167
8.3334 | -91.9165
8.3338 | -91.9165
8.3341 | -91.9159
8.3345 | -91.9155
8.3348 | -91.9157
8.3351 | -91.9159
8.3355 | -91.9156
8.3358 | -91.9149
8.3362 | -91.9143
8.3365 | -91.9141
8.3369 | -91.9137
8.3372 | -91.9137
8.3375 | -91.9133
8.3379 | -91.9136
8.3382 | -91.9134
8.3386 | -91.9131
8.3389 | -91.9124
8.3392 | -91.9119
8.3396 | -91.9118
8.3399 | -91.9116
8.3403 | -91.9119
8.3406 | -91.9118
8.3409 | -91.9116
8.3413 | -91.9113
8.3416 | -91.9109
8.342 | -91.9104
8.3423 | -91.9101
8.3426 | -91.9102
8.343 | -91.9101
8.3433 | -91.9087
8.3437 | -91.9073
8.344 | -91.907
8.3443 | -91.9069
8.3447 | -91.9069
8.345 | -91.9067
8.3454 | -91.9065
8.3457 | -91.9061
8.346 | -91.9062
8.3464 | -91.9058
8.3467 | -91.9052
8.3471 | -91.905
8.3474 | -91.9048
8.3477 | -91.9052
8.3481 | -91.9056
8.3484 | -91.9052
8.3488 | -91.9042
8.3491 | -91.9041
8.3494 | -91.9041
8.3498 | -91.9044
8.3501 | -91.9048
8.3505 | -91.905
8.3508 | -91.9049
8.3512 | -91.9047
8.3515 | -91.9045
8.3518 | -91.9041
8.3522 | -91.9032
8.3525 | -91.9027
8.3529 | -91.9029
8.3532 | -91.9025
8.3535 | -91.9012
8.3539 | -91.9008
8.3542 | -91.9011
8.3546 | -91.9016
8.3549 | -91.9014
8.3552 | -91.9014
8.3556 | -91.9019
8.3559 | -91.9021
8.3563 | -91.902
8.3566 | -91.902
8.3569 | -91.9016
8.3573 | -91.9008
8.3576 | -91.9006
8.358 | -91.9003
8.3583 | -91.8995
8.3586 | -91.8985
8.359 | -91.8978
8.3593 | -91.8977
8.3597 | -91.8975
8.36 | -91.8976
8.3603 | -91.8974
8.3607 | -91.8971
8.361 | -91.8971
8.3614 | -91.8971
8.3617 | -91.8967
8.362 | -91.8965
8.3624 | -91.8968
8.3627 | -91.8973
8.3631 | -91.8968
8.3634 | -91.8962
8.3637 | -91.896
8.3641 | -91.8944
8.3644 | -91.8939
8.3648 | -91.8938
8.3651 | -91.8932
8.3655 | -91.8933
8.3658 | -91.893
8.3661 | -91.8925
8.3665 | -91.8923
8.3668 | -91.8921
8.3672 | -91.8926
8.3675 | -91.8936
8.3678 | -91.8934
8.3682 | -91.8928
8.3685 | -91.8926
8.3689 | -91.8925
8.3692 | -91.8927
8.3695 | -91.8925
8.3699 | -91.8927
8.3702 | -91.8924
8.3706 | -91.8921
8.3709 | -91.8921
8.3712 | -91.8919
8.3716 | -91.8918
8.3719 | -91.8919
8.3723 | -91.8933
8.3726 | -91.8948
8.3729 | -91.895
8.3733 | -91.8945
8.3736 | -91.8942
8.374 | -91.8946
8.3743 | -91.8945
8.3746 | -91.8943
8.375 | -91.8942
8.3753 | -91.8941
8.3757 | -91.8944
8.376 | -91.8943
8.3763 | -91.8939
8.3767 | -91.8926
8.377 | -91.8922
8.3774 | -91.8921
8.3777 | -91.892
8.3781 | -91.8919
8.3784 | -91.8919
8.3787 | -91.892
8.3791 | -91.8917
8.3794 | -91.8914
8.3798 | -91.8916
8.3801 | -91.8921
8.3804 | -91.8925
8.3808 | -91.8925
8.3811 | -91.8924
8.3815 | -91.8924
8.3818 | -91.8925
8.3821 | -91.8926
8.3825 | -91.893
8.3828 | -91.8931
8.3832 | -91.893
8.3835 | -91.8926
8.3838 | -91.8926
8.3842 | -91.8929
8.3845 | -91.893
8.3849 | -91.893
8.3852 | -91.8941
8.3855 | -91.8944
8.3859 | -91.894
8.3862 | -91.8927
8.3866 | -91.8915
8.3869 | -91.8909
8.3872 | -91.8914
8.3876 | -91.892
8.3879 | -91.8926
8.3883 | -91.8927
8.3886 | -91.8927
8.3889 | -91.8943
8.3893 | -91.895
8.3896 | -91.8961
8.39 | -91.8964
8.3903 | -91.896
8.3906 | -91.8955
8.391 | -91.8952
8.3913 | -91.8954
8.3917 | -91.8954
8.392 | -91.8953
8.3924 | -91.8955
8.3927 | -91.8954
8.393 | -91.8953
8.3934 | -91.8955
8.3937 | -91.8947
8.3941 | -91.8931
8.3944 | -91.8919
8.3947 | -91.8918
8.3951 | -91.8927
8.3954 | -91.8919
8.3958 | -91.8916
8.3961 | -91.8916
8.3964 | -91.8916
8.3968 | -91.8913
8.3971 | -91.8913
8.3975 | -91.8912
8.3978 | -91.8914
8.3981 | -91.8909
8.3985 | -91.8905
8.3988 | -91.8903
8.3992 | -91.891
8.3995 | -91.8901
8.3998 | -91.8891
8.4002 | -91.8882
8.4005 | -91.888
8.4009 | -91.888
8.4012 | -91.8878
8.4015 | -91.8875
8.4019 | -91.8873
8.4022 | -91.8872
8.4026 | -91.8872
8.4029 | -91.887
8.4032 | -91.8864
8.4036 | -91.8863
8.4039 | -91.8868
8.4043 | -91.8871
8.4046 | -91.8876
8.4049 | -91.8873
8.4053 | -91.8866
8.4056 | -91.8865
8.406 | -91.8866
8.4063 | -91.8868
8.4067 | -91.8866
8.407 | -91.8865
8.4073 | -91.8866
8.4077 | -91.8863
8.408 | -91.8862
8.4084 | -91.8866
8.4087 | -91.8868
8.409 | -91.8867
8.4094 | -91.8861
8.4097 | -91.8861
8.4101 | -91.8863
8.4104 | -91.8867
8.4107 | -91.8867
8.4111 | -91.886
8.4114 | -91.8858
8.4118 | -91.8864
8.4121 | -91.887
8.4124 | -91.8864
8.4128 | -91.8861
8.4131 | -91.8859
8.4135 | -91.8859
8.4138 | -91.8865
8.4141 | -91.8865
8.4145 | -91.8865
8.4148 | -91.8865
8.4152 | -91.8874
8.4155 | -91.8878
8.4158 | -91.8879
8.4162 | -91.8882
8.4165 | -91.8884
8.4169 | -91.8885
8.4172 | -91.8886
8.4175 | -91.8885
8.4179 | -91.8881
8.4182 | -91.8879
8.4186 | -91.8882
8.4189 | -91.8886
8.4192 | -91.8884
8.4196 | -91.8882
8.4199 | -91.8883
8.4203 | -91.8881
8.4206 | -91.8882
8.421 | -91.8884
8.4213 | -91.8883
8.4216 | -91.8884
8.422 | -91.8884
8.4223 | -91.8886
8.4227 | -91.8888
8.423 | -91.8891
8.4233 | -91.8896
8.4237 | -91.889
8.424 | -91.889
8.4244 | -91.8892
8.4247 | -91.8892
8.425 | -91.8889
8.4254 | -91.8882
8.4257 | -91.8881
8.4261 | -91.8886
8.4264 | -91.8897
8.4267 | -91.8898
8.4271 | -91.8898
8.4274 | -91.8899
8.4278 | -91.89
8.4281 | -91.8898
8.4284 | -91.8899
8.4288 | -91.8904
8.4291 | -91.8903
8.4295 | -91.8905
8.4298 | -91.8901
8.4301 | -91.8894
8.4305 | -91.8892
8.4308 | -91.8896
8.4312 | -91.8897
8.4315 | -91.8897
8.4318 | -91.8893
8.4322 | -91.8885
8.4325 | -91.8885
8.4329 | -91.8886
8.4332 | -91.8888
8.4336 | -91.8881
8.4339 | -91.8876
8.4342 | -91.8877
8.4346 | -91.8879
8.4349 | -91.8875
8.4353 | -91.8872
8.4356 | -91.8874
8.4359 | -91.8875
8.4363 | -91.8872
8.4366 | -91.8868
8.437 | -91.8861
8.4373 | -91.8857
8.4376 | -91.8859
8.438 | -91.8861
8.4383 | -91.8863
8.4387 | -91.8867
8.439 | -91.887
8.4393 | -91.8872
8.4397 | -91.8875
8.44 | -91.8875
8.4404 | -91.888
8.4407 | -91.8888
8.441 | -91.8892
8.4414 | -91.8891
8.4417 | -91.8889
8.4421 | -91.889
8.4424 | -91.8895
8.4427 | -91.8896
8.4431 | -91.8894
8.4434 | -91.8888
8.4438 | -91.8887
8.4441 | -91.889
8.4444 | -91.8892
8.4448 | -91.8896
8.4451 | -91.8906
8.4455 | -91.8915
8.4458 | -91.8917
8.4461 | -91.8913
8.4465 | -91.8909
8.4468 | -91.8907
8.4472 | -91.8903
8.4475 | -91.8904
8.4479 | -91.8905
8.4482 | -91.8907
8.4485 | -91.8907
8.4489 | -91.8906
8.4492 | -91.8907
8.4496 | -91.8905
8.4499 | -91.8901
8.4502 | -91.8898
8.4506 | -91.8895
8.4509 | -91.8899
8.4513 | -91.8902
8.4516 | -91.8903
8.4519 | -91.8913
8.4523 | -91.8919
8.4526 | -91.8925
8.453 | -91.8928
8.4533 | -91.8927
8.4536 | -91.8929
8.454 | -91.8933
8.4543 | -91.8937
8.4547 | -91.8942
8.455 | -91.8947
8.4553 | -91.895
8.4557 | -91.895
8.456 | -91.8939
8.4564 | -91.8922
8.4567 | -91.8919
8.457 | -91.8921
8.4574 | -91.8925
8.4577 | -91.8931
8.4581 | -91.8934
8.4584 | -91.8929
8.4587 | -91.8923
8.4591 | -91.8919
8.4594 | -91.8922
8.4598 | -91.8923
8.4601 | -91.8919
8.4604 | -91.8924
8.4608 | -91.893
8.4611 | -91.8925
8.4615 | -91.8921
8.4618 | -91.893
8.4622 | -91.8937
8.4625 | -91.8939
8.4628 | -91.8941
8.4632 | -91.8943
8.4635 | -91.8945
8.4639 | -91.8948
8.4642 | -91.8948
8.4645 | -91.8948
8.4649 | -91.8949
8.4652 | -91.8949
8.4656 | -91.8953
8.4659 | -91.8958
8.4662 | -91.8967
8.4666 | -91.8969
8.4669 | -91.897
8.4673 | -91.8972
8.4676 | -91.8982
8.4679 | -91.8988
8.4683 | -91.9002
8.4686 | -91.9005
8.469 | -91.9009
8.4693 | -91.9014
8.4696 | -91.9015
8.47 | -91.9013
8.4703 | -91.9011
8.4707 | -91.9013
8.471 | -91.9015
8.4713 | -91.9006
8.4717 | -91.8999
8.472 | -91.9002
8.4724 | -91.9006
8.4727 | -91.9008
8.473 | -91.9009
8.4734 | -91.9014
8.4737 | -91.9019
8.4741 | -91.9022
8.4744 | -91.9025
8.4747 | -91.9026
8.4751 | -91.9022
8.4754 | -91.9016
8.4758 | -91.9017
8.4761 | -91.902
8.4765 | -91.9022
8.4768 | -91.9025
8.4771 | -91.9027
8.4775 | -91.9028
8.4778 | -91.9016
8.4782 | -91.9007
8.4785 | -91.9005
8.4788 | -91.9006
8.4792 | -91.9007
8.4795 | -91.9007
8.4799 | -91.9
8.4802 | -91.8994
8.4805 | -91.899
8.4809 | -91.899
8.4812 | -91.8986
8.4816 | -91.8992
8.4819 | -91.8994
8.4822 | -91.8989
8.4826 | -91.8992
8.4829 | -91.8994
8.4833 | -91.8996
8.4836 | -91.8998
8.4839 | -91.9001
8.4843 | -91.9003
8.4846 | -91.9002
8.485 | -91.9002
8.4853 | -91.9006
8.4856 | -91.9005
8.486 | -91.9007
8.4863 | -91.901
8.4867 | -91.9012
8.487 | -91.9011
8.4873 | -91.9012
8.4877 | -91.901
8.488 | -91.9012
8.4884 | -91.9011
8.4887 | -91.9013
8.4891 | -91.9014
8.4894 | -91.9012
8.4897 | -91.9011
8.4901 | -91.9013
8.4904 | -91.9016
8.4908 | -91.9022
8.4911 | -91.9023
8.4914 | -91.9022
8.4918 | -91.9023
8.4921 | -91.9026
8.4925 | -91.9031
8.4928 | -91.9027
8.4931 | -91.9029
8.4935 | -91.9031
8.4938 | -91.9033
8.4942 | -91.9032
8.4945 | -91.9033
8.4948 | -91.9034
8.4952 | -91.9035
8.4955 | -91.9036
8.4959 | -91.9036
8.4962 | -91.9039
8.4965 | -91.904
8.4969 | -91.9042
8.4972 | -91.9046
8.4976 | -91.9043
8.4979 | -91.9049
8.4982 | -91.9053
8.4986 | -91.9055
8.4989 | -91.9051
8.4993 | -91.9041
8.4996 | -91.904
8.4999 | -91.9044
8.5003 | -91.9049
8.5006 | -91.9056
8.501 | -91.9059
8.5013 | -91.906
8.5016 | -91.9062
8.502 | -91.9063
8.5023 | -91.9063
8.5027 | -91.9064
8.503 | -91.907
8.5033 | -91.9078
8.5037 | -91.9091
8.504 | -91.9106
8.5044 | -91.9113
8.5047 | -91.9116
8.5051 | -91.9111
8.5054 | -91.9107
8.5057 | -91.9104
8.5061 | -91.9101
8.5064 | -91.9093
8.5068 | -91.9082
8.5071 | -91.9084
8.5074 | -91.9085
8.5078 | -91.9079
8.5081 | -91.9077
8.5085 | -91.9079
8.5088 | -91.9078
8.5091 | -91.9077
8.5095 | -91.9076
8.5098 | -91.9076
8.5102 | -91.9078
8.5105 | -91.9078
8.5108 | -91.908
8.5112 | -91.9078
8.5115 | -91.9076
8.5119 | -91.9078
8.5122 | -91.908
8.5125 | -91.9081
8.5129 | -91.9081
8.5132 | -91.9081
8.5136 | -91.9079
8.5139 | -91.9073
8.5142 | -91.9071
8.5146 | -91.9068
8.5149 | -91.9065
8.5153 | -91.9057
8.5156 | -91.9052
8.5159 | -91.9054
8.5163 | -91.9052
8.5166 | -91.9052
8.517 | -91.9052
8.5173 | -91.9056
8.5177 | -91.9059
8.518 | -91.9062
8.5183 | -91.9063
8.5187 | -91.9065
8.519 | -91.9068
8.5194 | -91.9064
8.5197 | -91.9063
8.52 | -91.9065
8.5204 | -91.9068
8.5207 | -91.9072
8.5211 | -91.9072
8.5214 | -91.907
8.5217 | -91.9069
8.5221 | -91.9067
8.5224 | -91.9065
8.5228 | -91.9065
8.5231 | -91.9061
8.5234 | -91.9064
8.5238 | -91.9062
8.5241 | -91.9056
8.5245 | -91.9054
8.5248 | -91.9058
8.5251 | -91.9063
8.5255 | -91.906
8.5258 | -91.9056
8.5262 | -91.9055
8.5265 | -91.9057
8.5268 | -91.9059
8.5272 | -91.9057
8.5275 | -91.906
8.5279 | -91.906
8.5282 | -91.9057
8.5285 | -91.9057
8.5289 | -91.906
8.5292 | -91.9062
8.5296 | -91.9065
8.5299 | -91.9069
8.5302 | -91.9069
8.5306 | -91.9066
8.5309 | -91.9067
8.5313 | -91.9064
8.5316 | -91.9056
8.532 | -91.9056
8.5323 | -91.9059
8.5326 | -91.9059
8.533 | -91.9062
8.5333 | -91.9059
8.5337 | -91.9046
8.534 | -91.9034
8.5343 | -91.9032
8.5347 | -91.9034
8.535 | -91.9033
8.5354 | -91.9034
8.5357 | -91.9032
8.536 | -91.9031
8.5364 | -91.9033
8.5367 | -91.9031
8.5371 | -91.9032
8.5374 | -91.9031
8.5377 | -91.9023
8.5381 | -91.9016
8.5384 | -91.9018
8.5388 | -91.9019
8.5391 | -91.9017
8.5394 | -91.9019
8.5398 | -91.9025
8.5401 | -91.9026
8.5405 | -91.9021
8.5408 | -91.9021
8.5411 | -91.9019
8.5415 | -91.9018
8.5418 | -91.9014
8.5422 | -91.9018
8.5425 | -91.9026
8.5428 | -91.9027
8.5432 | -91.9025
8.5435 | -91.9022
8.5439 | -91.9021
8.5442 | -91.9021
8.5445 | -91.9022
8.5449 | -91.9021
8.5452 | -91.9019
8.5456 | -91.902
8.5459 | -91.9021
8.5463 | -91.9021
8.5466 | -91.9019
8.5469 | -91.9022
8.5473 | -91.9021
8.5476 | -91.9021
8.548 | -91.9022
8.5483 | -91.9023
8.5486 | -91.9023
8.549 | -91.9022
8.5493 | -91.902
8.5497 | -91.9016
8.55 | -91.9012
8.5503 | -91.9007
8.5507 | -91.9001
8.551 | -91.8998
8.5514 | -91.9
8.5517 | -91.9003
8.552 | -91.9
8.5524 | -91.9
8.5527 | -91.8999
8.5531 | -91.9
8.5534 | -91.9001
8.5537 | -91.9
8.5541 | -91.8998
8.5544 | -91.8999
8.5548 | -91.8994
8.5551 | -91.8992
8.5554 | -91.8993
8.5558 | -91.8995
8.5561 | -91.8991
8.5565 | -91.899
8.5568 | -91.8988
8.5571 | -91.899
8.5575 | -91.8992
8.5578 | -91.8993
8.5582 | -91.8986
8.5585 | -91.898
8.5588 | -91.8979
8.5592 | -91.8979
8.5595 | -91.8979
8.5599 | -91.8976
8.5602 | -91.8968
8.5606 | -91.8962
8.5609 | -91.8959
8.5612 | -91.8957
8.5616 | -91.8948
8.5619 | -91.894
8.5623 | -91.8934
8.5626 | -91.8932
8.5629 | -91.8928
8.5633 | -91.8923
8.5636 | -91.8923
8.564 | -91.8922
8.5643 | -91.8921
8.5646 | -91.892
8.565 | -91.8919
8.5653 | -91.8916
8.5657 | -91.8911
8.566 | -91.8912
8.5663 | -91.8911
8.5667 | -91.8913
8.567 | -91.8915
8.5674 | -91.8914
8.5677 | -91.8913
8.568 | -91.8918
8.5684 | -91.8925
8.5687 | -91.8929
8.5691 | -91.8924
8.5694 | -91.892
8.5697 | -91.8919
8.5701 | -91.892
8.5704 | -91.8921
8.5708 | -91.8921
8.5711 | -91.8921
8.5714 | -91.8921
8.5718 | -91.8921
8.5721 | -91.8922
8.5725 | -91.8925
8.5728 | -91.8925
8.5732 | -91.8926
8.5735 | -91.893
8.5738 | -91.893
8.5742 | -91.8922
8.5745 | -91.8924
8.5749 | -91.8924
8.5752 | -91.8922
8.5755 | -91.892
8.5759 | -91.8919
8.5762 | -91.892
8.5766 | -91.8917
8.5769 | -91.891
8.5772 | -91.8907
8.5776 | -91.8903
8.5779 | -91.8902
8.5783 | -91.8902
8.5786 | -91.8898
8.5789 | -91.8894
8.5793 | -91.889
8.5796 | -91.8885
8.58 | -91.8884
8.5803 | -91.8883
8.5806 | -91.8881
8.581 | -91.8872
8.5813 | -91.8868
8.5817 | -91.8862
8.582 | -91.8857
8.5823 | -91.8856
8.5827 | -91.8854
8.583 | -91.8853
8.5834 | -91.8848
8.5837 | -91.8845
8.584 | -91.8836
8.5844 | -91.8825
8.5847 | -91.8813
8.5851 | -91.8805
8.5854 | -91.8802
8.5857 | -91.8798
8.5861 | -91.8796
8.5864 | -91.8797
8.5868 | -91.8794
8.5871 | -91.879
8.5875 | -91.879
8.5878 | -91.8791
8.5881 | -91.879
8.5885 | -91.8784
8.5888 | -91.8774
8.5892 | -91.8773
8.5895 | -91.877
8.5898 | -91.8761
8.5902 | -91.876
8.5905 | -91.8764
8.5909 | -91.8765
8.5912 | -91.8762
8.5915 | -91.8756
8.5919 | -91.8755
8.5922 | -91.8754
8.5926 | -91.8751
8.5929 | -91.8749
8.5932 | -91.8749
8.5936 | -91.8747
8.5939 | -91.8743
8.5943 | -91.8742
8.5946 | -91.874
8.5949 | -91.8736
8.5953 | -91.8731
8.5956 | -91.8728
8.596 | -91.8723
8.5963 | -91.8715
8.5966 | -91.8711
8.597 | -91.8708
8.5973 | -91.8705
8.5977 | -91.8709
8.598 | -91.8715
8.5983 | -91.8721
8.5987 | -91.873
8.599 | -91.8729
8.5994 | -91.8725
8.5997 | -91.8729
8.6 | -91.8724
8.6004 | -91.8718
8.6007 | -91.8716
8.6011 | -91.8717
8.6014 | -91.8716
8.6018 | -91.871
8.6021 | -91.87
8.6024 | -91.8684
8.6028 | -91.8674
8.6031 | -91.867
8.6035 | -91.8668
8.6038 | -91.8665
8.6041 | -91.8665
8.6045 | -91.8675
8.6048 | -91.8684
8.6052 | -91.8683
8.6055 | -91.8678
8.6058 | -91.8667
8.6062 | -91.8665
8.6065 | -91.8664
8.6069 | -91.8658
8.6072 | -91.8653
8.6075 | -91.865
8.6079 | -91.8651
8.6082 | -91.8652
8.6086 | -91.8648
8.6089 | -91.865
8.6092 | -91.8657
8.6096 | -91.8651
8.6099 | -91.8644
8.6103 | -91.8647
8.6106 | -91.8642
8.6109 | -91.8637
8.6113 | -91.8639
8.6116 | -91.8636
8.612 | -91.8629
8.6123 | -91.8623
8.6126 | -91.862
8.613 | -91.8619
8.6133 | -91.8612
8.6137 | -91.8606
8.614 | -91.8603
8.6143 | -91.8601
8.6147 | -91.8599
8.615 | -91.8594
8.6154 | -91.8595
8.6157 | -91.859
8.6161 | -91.8584
8.6164 | -91.858
8.6167 | -91.8577
8.6171 | -91.8569
8.6174 | -91.8562
8.6178 | -91.8561
8.6181 | -91.8559
8.6184 | -91.8548
8.6188 | -91.8532
8.6191 | -91.8529
8.6195 | -91.8527
8.6198 | -91.8524
8.6201 | -91.8521
8.6205 | -91.8514
8.6208 | -91.851
8.6212 | -91.8508
8.6215 | -91.8505
8.6218 | -91.8501
8.6222 | -91.8502
8.6225 | -91.8496
8.6229 | -91.8484
8.6232 | -91.8472
8.6235 | -91.8469
8.6239 | -91.8466
8.6242 | -91.8462
8.6246 | -91.8459
8.6249 | -91.8456
8.6252 | -91.8455
8.6256 | -91.8448
8.6259 | -91.8445
8.6263 | -91.8443
8.6266 | -91.8439
8.6269 | -91.8436
8.6273 | -91.8433
8.6276 | -91.8425
8.628 | -91.842
8.6283 | -91.8419
8.6287 | -91.8412
8.629 | -91.8407
8.6293 | -91.8405
8.6297 | -91.8401
8.63 | -91.8407
8.6304 | -91.8417
8.6307 | -91.8425
8.631 | -91.8428
8.6314 | -91.8425
8.6317 | -91.8424
8.6321 | -91.8422
8.6324 | -91.8419
8.6327 | -91.8413
8.6331 | -91.8409
8.6334 | -91.8399
8.6338 | -91.8389
8.6341 | -91.8384
8.6344 | -91.8379
8.6348 | -91.838
8.6351 | -91.8378
8.6355 | -91.8381
8.6358 | -91.8382
8.6361 | -91.8376
8.6365 | -91.8371
8.6368 | -91.8372
8.6372 | -91.8372
8.6375 | -91.8371
8.6378 | -91.837
8.6382 | -91.8367
8.6385 | -91.8366
8.6389 | -91.8366
8.6392 | -91.8365
8.6395 | -91.8376
8.6399 | -91.8382
8.6402 | -91.838
8.6406 | -91.8378
8.6409 | -91.8374
8.6412 | -91.8371
8.6416 | -91.8368
8.6419 | -91.8365
8.6423 | -91.8357
8.6426 | -91.8347
8.643 | -91.8342
8.6433 | -91.8341
8.6436 | -91.8338
8.644 | -91.8336
8.6443 | -91.8333
8.6447 | -91.833
8.645 | -91.8328
8.6453 | -91.8327
8.6457 | -91.8334
8.646 | -91.8335
8.6464 | -91.8331
8.6467 | -91.8323
8.647 | -91.8311
8.6474 | -91.8301
8.6477 | -91.8296
8.6481 | -91.8302
8.6484 | -91.8306
8.6487 | -91.8316
8.6491 | -91.8323
8.6494 | -91.8322
8.6498 | -91.8315
8.6501 | -91.8316
8.6504 | -91.8316
8.6508 | -91.8319
8.6511 | -91.832
8.6515 | -91.8317
8.6518 | -91.8313
8.6521 | -91.8308
8.6525 | -91.8303
8.6528 | -91.8298
8.6532 | -91.8293
8.6535 | -91.8289
8.6538 | -91.8284
8.6542 | -91.8278
8.6545 | -91.8273
8.6549 | -91.8273
8.6552 | -91.827
8.6555 | -91.8265
8.6559 | -91.8259
8.6562 | -91.8257
8.6566 | -91.8254
8.6569 | -91.8259
8.6573 | -91.8256
8.6576 | -91.8254
8.6579 | -91.825
8.6583 | -91.8244
8.6586 | -91.8243
8.659 | -91.8236
8.6593 | -91.823
8.6596 | -91.8226
8.66 | -91.8225
8.6603 | -91.8222
8.6607 | -91.8216
8.661 | -91.8216
8.6613 | -91.8217
8.6617 | -91.8213
8.662 | -91.8209
8.6624 | -91.8207
8.6627 | -91.8201
8.663 | -91.8198
8.6634 | -91.8199
8.6637 | -91.8199
8.6641 | -91.8199
8.6644 | -91.8198
8.6647 | -91.8195
8.6651 | -91.8192
8.6654 | -91.819
8.6658 | -91.8184
8.6661 | -91.8177
8.6664 | -91.8175
8.6668 | -91.8171
8.6671 | -91.8168
8.6675 | -91.8166
8.6678 | -91.8162
8.6681 | -91.8159
8.6685 | -91.8158
8.6688 | -91.8155
8.6692 | -91.8154
8.6695 | -91.8153
8.6698 | -91.8151
8.6702 | -91.8149
8.6705 | -91.8144
8.6709 | -91.8142
8.6712 | -91.8141
8.6716 | -91.8133
8.6719 | -91.8124
8.6722 | -91.8119
8.6726 | -91.8116
8.6729 | -91.8108
8.6733 | -91.8101
8.6736 | -91.8098
8.6739 | -91.8095
8.6743 | -91.8097
8.6746 | -91.8094
8.675 | -91.8089
8.6753 | -91.8086
8.6756 | -91.8083
8.676 | -91.8085
8.6763 | -91.8082
8.6767 | -91.8082
8.677 | -91.808
8.6773 | -91.8079
8.6777 | -91.8075
8.678 | -91.8071
8.6784 | -91.8068
8.6787 | -91.8067
8.679 | -91.8066
8.6794 | -91.8066
8.6797 | -91.8068
8.6801 | -91.8066
8.6804 | -91.8064
8.6807 | -91.8061
8.6811 | -91.8057
8.6814 | -91.8055
8.6818 | -91.8057
8.6821 | -91.8059
8.6824 | -91.8058
8.6828 | -91.8055
8.6831 | -91.805
8.6835 | -91.8048
8.6838 | -91.8047
8.6842 | -91.8051
8.6845 | -91.8052
8.6848 | -91.8052
8.6852 | -91.8049
8.6855 | -91.8044
8.6859 | -91.8038
8.6862 | -91.8036
8.6865 | -91.8033
8.6869 | -91.8029
8.6872 | -91.8027
8.6876 | -91.8027
8.6879 | -91.8024
8.6882 | -91.8022
8.6886 | -91.8025
8.6889 | -91.8023
8.6893 | -91.8025
8.6896 | -91.8024
8.6899 | -91.8019
8.6903 | -91.8026
8.6906 | -91.8031
8.691 | -91.8028
8.6913 | -91.8022
8.6916 | -91.8022
8.692 | -91.8022
8.6923 | -91.802
8.6927 | -91.8022
8.693 | -91.8025
8.6933 | -91.8028
8.6937 | -91.8026
8.694 | -91.803
8.6944 | -91.803
8.6947 | -91.8033
8.695 | -91.8035
8.6954 | -91.8033
8.6957 | -91.8034
8.6961 | -91.8036
8.6964 | -91.8032
8.6967 | -91.8025
8.6971 | -91.8022
8.6974 | -91.802
8.6978 | -91.8019
8.6981 | -91.8017
8.6984 | -91.8017
8.6988 | -91.8024
8.6991 | -91.8025
8.6995 | -91.8026
8.6998 | -91.8025
8.7002 | -91.8021
8.7005 | -91.8015
8.7008 | -91.8013
8.7012 | -91.8013
8.7015 | -91.8021
8.7019 | -91.8022
8.7022 | -91.8021
8.7025 | -91.8019
8.7029 | -91.8018
8.7032 | -91.8016
8.7036 | -91.8011
8.7039 | -91.8015
8.7042 | -91.801
8.7046 | -91.8008
8.7049 | -91.8006
8.7053 | -91.8003
8.7056 | -91.8002
8.7059 | -91.8
8.7063 | -91.8
8.7066 | -91.7999
8.707 | -91.8
8.7073 | -91.8001
8.7076 | -91.8006
8.708 | -91.8005
8.7083 | -91.8005
8.7087 | -91.801
8.709 | -91.801
8.7093 | -91.8011
8.7097 | -91.8014
8.71 | -91.8017
8.7104 | -91.8016
8.7107 | -91.8016
8.711 | -91.8018
8.7114 | -91.8037
8.7117 | -91.8043
8.7121 | -91.8041
8.7124 | -91.8036
8.7128 | -91.8041
8.7131 | -91.8042
8.7134 | -91.8046
8.7138 | -91.8053
8.7141 | -91.8058
8.7145 | -91.8061
8.7148 | -91.8065
8.7151 | -91.807
8.7155 | -91.8077
8.7158 | -91.808
8.7162 | -91.8079
8.7165 | -91.8091
8.7168 | -91.8101
8.7172 | -91.8112
8.7175 | -91.8115
8.7179 | -91.8116
8.7182 | -91.8117
8.7185 | -91.812
8.7189 | -91.813
8.7192 | -91.8132
8.7196 | -91.8145
8.7199 | -91.8145
8.7202 | -91.8151
8.7206 | -91.8151
8.7209 | -91.8161
8.7213 | -91.8172
8.7216 | -91.8171
8.7219 | -91.817
8.7223 | -91.8186
8.7226 | -91.8208
8.723 | -91.8217
8.7233 | -91.8224
8.7236 | -91.8235
8.724 | -91.8257
8.7243 | -91.8267
8.7247 | -91.8276
8.725 | -91.8293
8.7253 | -91.8304
8.7257 | -91.8318
8.726 | -91.8323
8.7264 | -91.8336
8.7267 | -91.8338
8.7271 | -91.8345
8.7274 | -91.835
8.7277 | -91.8352
8.7281 | -91.8352
8.7284 | -91.8349
8.7288 | -91.8342
8.7291 | -91.8341
8.7294 | -91.8337
8.7298 | -91.8338
8.7301 | -91.8335
8.7305 | -91.8334
8.7308 | -91.8335
8.7311 | -91.8336
8.7315 | -91.8335
8.7318 | -91.8335
8.7322 | -91.8333
8.7325 | -91.833
8.7328 | -91.8326
8.7332 | -91.8327
8.7335 | -91.8323
8.7339 | -91.8323
8.7342 | -91.8329
8.7345 | -91.8341
8.7349 | -91.8344
8.7352 | -91.8343
8.7356 | -91.834
8.7359 | -91.8334
8.7362 | -91.8326
8.7366 | -91.8326
8.7369 | -91.8328
8.7373 | -91.8329
8.7376 | -91.8335
8.7379 | -91.834
8.7383 | -91.8345
8.7386 | -91.8346
8.739 | -91.8347
8.7393 | -91.8348
8.7396 | -91.8351
8.74 | -91.835
8.7403 | -91.8354
8.7407 | -91.8358
8.741 | -91.8361
8.7414 | -91.8366
8.7417 | -91.8369
8.742 | -91.8371
8.7424 | -91.8372
8.7427 | -91.8374
8.7431 | -91.8382
8.7434 | -91.8387
8.7437 | -91.8391
8.7441 | -91.8392
8.7444 | -91.8388
8.7448 | -91.8387
8.7451 | -91.8391
8.7454 | -91.8395
8.7458 | -91.8396
8.7461 | -91.8388
8.7465 | -91.8389
8.7468 | -91.8392
8.7471 | -91.8393
8.7475 | -91.8393
8.7478 | -91.8393
8.7482 | -91.8391
8.7485 | -91.8388
8.7488 | -91.839
8.7492 | -91.8391
8.7495 | -91.8396
8.7499 | -91.8397
8.7502 | -91.8395
8.7505 | -91.8401
8.7509 | -91.8409
8.7512 | -91.8411
8.7516 | -91.8414
8.7519 | -91.8418
8.7522 | -91.8419
8.7526 | -91.8422
8.7529 | -91.8433
8.7533 | -91.8436
8.7536 | -91.8439
8.7539 | -91.8438
8.7543 | -91.8432
8.7546 | -91.8434
8.755 | -91.8438
8.7553 | -91.8441
8.7557 | -91.8444
8.756 | -91.8445
8.7563 | -91.8438
8.7567 | -91.8441
8.757 | -91.8446
8.7574 | -91.8448
8.7577 | -91.8454
8.758 | -91.8462
8.7584 | -91.8466
8.7587 | -91.8478
8.7591 | -91.8489
8.7594 | -91.8492
8.7597 | -91.8496
8.7601 | -91.8502
8.7604 | -91.8508
8.7608 | -91.8519
8.7611 | -91.8526
8.7614 | -91.8526
8.7618 | -91.8523
8.7621 | -91.8522
8.7625 | -91.8526
8.7628 | -91.8536
8.7631 | -91.8541
8.7635 | -91.8545
8.7638 | -91.855
8.7642 | -91.8554
8.7645 | -91.8557
8.7648 | -91.8555
8.7652 | -91.8549
8.7655 | -91.8547
8.7659 | -91.8545
8.7662 | -91.855
8.7665 | -91.8553
8.7669 | -91.8552
8.7672 | -91.855
8.7676 | -91.8552
8.7679 | -91.8557
8.7683 | -91.8562
8.7686 | -91.8565
8.7689 | -91.8575
8.7693 | -91.8579
8.7696 | -91.8585
8.77 | -91.8593
8.7703 | -91.8598
8.7706 | -91.8602
8.771 | -91.8615
8.7713 | -91.8627
8.7717 | -91.8632
8.772 | -91.8636
8.7723 | -91.8637
8.7727 | -91.8641
8.773 | -91.8646
8.7734 | -91.865
8.7737 | -91.8655
8.774 | -91.866
8.7744 | -91.8666
8.7747 | -91.867
8.7751 | -91.8677
8.7754 | -91.8682
8.7757 | -91.8686
8.7761 | -91.8691
8.7764 | -91.8696
8.7768 | -91.8698
8.7771 | -91.87
8.7774 | -91.8705
8.7778 | -91.8713
8.7781 | -91.8731
8.7785 | -91.8742
8.7788 | -91.8751
8.7791 | -91.8753
8.7795 | -91.876
8.7798 | -91.8767
8.7802 | -91.8772
8.7805 | -91.8774
8.7808 | -91.878
8.7812 | -91.8787
8.7815 | -91.8791
8.7819 | -91.8799
8.7822 | -91.8804
8.7826 | -91.8818
8.7829 | -91.8823
8.7832 | -91.8829
8.7836 | -91.8831
8.7839 | -91.8836
8.7843 | -91.884
8.7846 | -91.8848
8.7849 | -91.8851
8.7853 | -91.8855
8.7856 | -91.8863
8.786 | -91.8869
8.7863 | -91.8877
8.7866 | -91.8887
8.787 | -91.889
8.7873 | -91.8893
8.7877 | -91.8903
8.788 | -91.891
8.7883 | -91.8916
8.7887 | -91.8925
8.789 | -91.894
8.7894 | -91.8948
8.7897 | -91.8955
8.79 | -91.8961
8.7904 | -91.8971
8.7907 | -91.8979
8.7911 | -91.8984
8.7914 | -91.899
8.7917 | -91.8995
8.7921 | -91.9006
8.7924 | -91.9005
8.7928 | -91.9008
8.7931 | -91.9016
8.7934 | -91.9021
8.7938 | -91.9023
8.7941 | -91.9026
8.7945 | -91.9031
8.7948 | -91.9033
8.7951 | -91.9038
8.7955 | -91.9043
8.7958 | -91.9049
8.7962 | -91.9055
8.7965 | -91.9059
8.7969 | -91.9065
8.7972 | -91.9067
8.7975 | -91.9062
8.7979 | -91.9065
8.7982 | -91.9065
8.7986 | -91.9059
8.7989 | -91.9058
8.7992 | -91.9065
8.7996 | -91.9073
8.7999 | -91.9082
8.8003 | -91.9089
8.8006 | -91.9095
8.8009 | -91.91
8.8013 | -91.9109
8.8016 | -91.9125
8.802 | -91.9132
8.8023 | -91.9144
8.8026 | -91.9151
8.803 | -91.9158
8.8033 | -91.9168
8.8037 | -91.9174
8.804 | -91.9183
8.8043 | -91.9194
8.8047 | -91.9201
8.805 | -91.9209
8.8054 | -91.9216
8.8057 | -91.9225
8.806 | -91.9232
8.8064 | -91.9243
8.8067 | -91.9256
8.8071 | -91.926
8.8074 | -91.9263
8.8077 | -91.927
8.8081 | -91.9278
8.8084 | -91.9284
8.8088 | -91.9296
8.8091 | -91.9302
8.8094 | -91.9309
8.8098 | -91.9316
8.8101 | -91.9324
8.8105 | -91.9333
8.8108 | -91.934
8.8112 | -91.9346
8.8115 | -91.9355
8.8118 | -91.9362
8.8122 | -91.9372
8.8125 | -91.9381
8.8129 | -91.9388
8.8132 | -91.9396
8.8135 | -91.9403
8.8139 | -91.941
8.8142 | -91.9418
8.8146 | -91.9428
8.8149 | -91.9436
8.8152 | -91.9442
8.8156 | -91.9447
8.8159 | -91.9454
8.8163 | -91.9465
8.8166 | -91.9472
8.8169 | -91.9489
8.8173 | -91.9504
8.8176 | -91.9512
8.818 | -91.9524
8.8183 | -91.9539
8.8186 | -91.9545
8.819 | -91.9554
8.8193 | -91.9562
8.8197 | -91.9571
8.82 | -91.958
8.8203 | -91.9585
8.8207 | -91.959
8.821 | -91.9601
8.8214 | -91.9608
8.8217 | -91.9616
8.822 | -91.9628
8.8224 | -91.9635
8.8227 | -91.9641
8.8231 | -91.9648
8.8234 | -91.9655
8.8238 | -91.9664
8.8241 | -91.9672
8.8244 | -91.968
8.8248 | -91.969
8.8251 | -91.9703
8.8255 | -91.9715
8.8258 | -91.9723
8.8261 | -91.9734
8.8265 | -91.9745
8.8268 | -91.9757
8.8272 | -91.9765
8.8275 | -91.977
8.8278 | -91.978
8.8282 | -91.979
8.8285 | -91.9799
8.8289 | -91.9809
8.8292 | -91.9827
8.8295 | -91.9841
8.8299 | -91.9851
8.8302 | -91.9863
8.8306 | -91.9872
8.8309 | -91.9882
8.8312 | -91.989
8.8316 | -91.9894
8.8319 | -91.9907
8.8323 | -91.9918
8.8326 | -91.9917
8.8329 | -91.9923
8.8333 | -91.9933
8.8336 | -91.9948
8.834 | -91.9961
8.8343 | -91.9965
8.8346 | -91.9971
8.835 | -91.9981
8.8353 | -91.9991
8.8357 | -92
8.836 | -92.0012
8.8363 | -92.0026
8.8367 | -92.0035
8.837 | -92.0044
8.8374 | -92.0053
8.8377 | -92.0062
8.8381 | -92.0072
8.8384 | -92.0083
8.8387 | -92.0094
8.8391 | -92.0104
8.8394 | -92.0111
8.8398 | -92.0113
8.8401 | -92.0111
8.8404 | -92.0109
8.8408 | -92.0113
8.8411 | -92.0123
8.8415 | -92.0134
8.8418 | -92.0144
8.8421 | -92.0154
8.8425 | -92.0164
8.8428 | -92.0175
8.8432 | -92.0183
8.8435 | -92.0196
8.8438 | -92.021
8.8442 | -92.0223
8.8445 | -92.0231
8.8449 | -92.0238
8.8452 | -92.0247
8.8455 | -92.026
8.8459 | -92.0276
8.8462 | -92.0286
8.8466 | -92.0295
8.8469 | -92.0304
8.8472 | -92.0315
8.8476 | -92.0325
8.8479 | -92.0333
8.8483 | -92.0338
8.8486 | -92.0341
8.8489 | -92.0347
8.8493 | -92.0361
8.8496 | -92.0374
8.85 | -92.0382
8.8503 | -92.0393
8.8506 | -92.0403
8.851 | -92.0409
8.8513 | -92.0419
8.8517 | -92.0429
8.852 | -92.044
8.8524 | -92.0454
8.8527 | -92.0465
8.853 | -92.0478
8.8534 | -92.0494
8.8537 | -92.0501
8.8541 | -92.051
8.8544 | -92.052
8.8547 | -92.0527
8.8551 | -92.0536
8.8554 | -92.0544
8.8558 | -92.0554
8.8561 | -92.0566
8.8564 | -92.0581
8.8568 | -92.0592
8.8571 | -92.0602
8.8575 | -92.0614
8.8578 | -92.0634
8.8581 | -92.0654
8.8585 | -92.0655
8.8588 | -92.0654
8.8592 | -92.066
8.8595 | -92.0667
8.8598 | -92.0673
8.8602 | -92.0682
8.8605 | -92.0692
8.8609 | -92.0703
8.8612 | -92.0709
8.8615 | -92.072
8.8619 | -92.073
8.8622 | -92.074
8.8626 | -92.0752
8.8629 | -92.0763
8.8632 | -92.0773
8.8636 | -92.0774
8.8639 | -92.0777
8.8643 | -92.0782
8.8646 | -92.0788
8.8649 | -92.0795
8.8653 | -92.0805
8.8656 | -92.0815
8.866 | -92.0828
8.8663 | -92.0844
8.8667 | -92.0857
8.867 | -92.087
8.8673 | -92.0887
8.8677 | -92.0898
8.868 | -92.0916
8.8684 | -92.0934
8.8687 | -92.0946
8.869 | -92.096
8.8694 | -92.0969
8.8697 | -92.098
8.8701 | -92.099
8.8704 | -92.0993
8.8707 | -92.1005
8.8711 | -92.1016
8.8714 | -92.1026
8.8718 | -92.1034
8.8721 | -92.1045
8.8724 | -92.106
8.8728 | -92.1068
8.8731 | -92.1072
8.8735 | -92.1082
8.8738 | -92.1092
8.8741 | -92.1102
8.8745 | -92.1114
8.8748 | -92.1127
8.8752 | -92.1133
8.8755 | -92.1137
8.8758 | -92.1146
8.8762 | -92.1153
8.8765 | -92.1163
8.8769 | -92.1173
8.8772 | -92.118
8.8775 | -92.1185
8.8779 | -92.1199
8.8782 | -92.121
8.8786 | -92.1223
8.8789 | -92.1233
8.8792 | -92.1239
8.8796 | -92.1255
8.8799 | -92.1261
8.8803 | -92.1268
8.8806 | -92.1276
8.881 | -92.1287
8.8813 | -92.1304
8.8816 | -92.1316
8.882 | -92.1326
8.8823 | -92.1335
8.8827 | -92.1343
8.883 | -92.1352
8.8833 | -92.1357
8.8837 | -92.1364
8.884 | -92.1373
8.8844 | -92.1387
8.8847 | -92.1397
8.885 | -92.1411
8.8854 | -92.1425
8.8857 | -92.1436
8.8861 | -92.1444
8.8864 | -92.1445
8.8867 | -92.145
8.8871 | -92.1461
8.8874 | -92.1471
8.8878 | -92.1481
8.8881 | -92.149
8.8884 | -92.15
8.8888 | -92.1514
8.8891 | -92.1521
8.8895 | -92.1532
8.8898 | -92.1545
8.8901 | -92.1561
8.8905 | -92.1575
8.8908 | -92.1588
8.8912 | -92.1599
8.8915 | -92.1609
8.8918 | -92.1619
8.8922 | -92.1629
8.8925 | -92.1642
8.8929 | -92.1659
8.8932 | -92.167
8.8935 | -92.1673
8.8939 | -92.1686
8.8942 | -92.1695
8.8946 | -92.1704
8.8949 | -92.1719
8.8953 | -92.1729
8.8956 | -92.1743
8.8959 | -92.1753
8.8963 | -92.176
8.8966 | -92.1768
8.897 | -92.1783
8.8973 | -92.1796
8.8976 | -92.1805
8.898 | -92.1818
8.8983 | -92.183
8.8987 | -92.1842
8.899 | -92.1857
8.8993 | -92.1867
8.8997 | -92.1876
8.9 | -92.1888
8.9004 | -92.1893
8.9007 | -92.1905
8.901 | -92.1916
8.9014 | -92.1929
8.9017 | -92.1944
8.9021 | -92.1961
8.9024 | -92.1976
8.9027 | -92.1983
8.9031 | -92.1992
8.9034 | -92.2005
8.9038 | -92.2021
8.9041 | -92.204
8.9044 | -92.2047
8.9048 | -92.2053
8.9051 | -92.206
8.9055 | -92.2065
8.9058 | -92.2076
8.9061 | -92.2086
8.9065 | -92.2096
8.9068 | -92.211
8.9072 | -92.2124
8.9075 | -92.2136
8.9078 | -92.2146
8.9082 | -92.2159
8.9085 | -92.2169
8.9089 | -92.2172
8.9092 | -92.217
8.9096 | -92.2176
8.9099 | -92.2187
8.9102 | -92.2199
8.9106 | -92.2207
8.9109 | -92.2223
8.9113 | -92.2237
8.9116 | -92.2247
8.9119 | -92.2257
8.9123 | -92.2264
8.9126 | -92.2279
8.913 | -92.2298
8.9133 | -92.2308
8.9136 | -92.2318
8.914 | -92.2327
8.9143 | -92.2338
8.9147 | -92.2349
8.915 | -92.2357
8.9153 | -92.2373
8.9157 | -92.2391
8.916 | -92.2403
8.9164 | -92.2414
8.9167 | -92.2426
8.917 | -92.2436
8.9174 | -92.2446
8.9177 | -92.2463
8.9181 | -92.2468
8.9184 | -92.2473
8.9187 | -92.2484
8.9191 | -92.2494
8.9194 | -92.2499
8.9198 | -92.2509
8.9201 | -92.2502
8.9204 | -92.2501
8.9208 | -92.2511
8.9211 | -92.2518
8.9215 | -92.253
8.9218 | -92.2538
8.9222 | -92.2546
8.9225 | -92.2555
8.9228 | -92.2556
8.9232 | -92.2558
8.9235 | -92.2568
8.9239 | -92.2576
8.9242 | -92.2585
8.9245 | -92.2595
8.9249 | -92.2605
8.9252 | -92.2611
8.9256 | -92.2619
8.9259 | -92.2637
8.9262 | -92.2657
8.9266 | -92.2667
8.9269 | -92.2677
8.9273 | -92.2684
8.9276 | -92.2692
8.9279 | -92.2703
8.9283 | -92.2719
8.9286 | -92.2728
8.929 | -92.2736
8.9293 | -92.2743
8.9296 | -92.2748
8.93 | -92.2759
8.9303 | -92.2772
8.9307 | -92.2785
8.931 | -92.2815
8.9313 | -92.2834
8.9317 | -92.2841
8.932 | -92.2854
8.9324 | -92.2867
8.9327 | -92.2875
8.933 | -92.2885
8.9334 | -92.2894
8.9337 | -92.2905
8.9341 | -92.2912
8.9344 | -92.2918
8.9347 | -92.2926
8.9351 | -92.2938
8.9354 | -92.2946
8.9358 | -92.2954
8.9361 | -92.2961
8.9365 | -92.2966
8.9368 | -92.2975
8.9371 | -92.2984
8.9375 | -92.2994
8.9378 | -92.3006
8.9382 | -92.3018
8.9385 | -92.3027
8.9388 | -92.3035
8.9392 | -92.3039
8.9395 | -92.3039
8.9399 | -92.3042
8.9402 | -92.3044
8.9405 | -92.3052
8.9409 | -92.3061
8.9412 | -92.3068
8.9416 | -92.3073
8.9419 | -92.3084
8.9422 | -92.3091
8.9426 | -92.3098
8.9429 | -92.3107
8.9433 | -92.3129
8.9436 | -92.315
8.9439 | -92.3157
8.9443 | -92.3163
8.9446 | -92.3172
8.945 | -92.318
8.9453 | -92.3183
8.9456 | -92.3184
8.946 | -92.3184
8.9463 | -92.3192
8.9467 | -92.3205
8.947 | -92.3216
8.9473 | -92.3223
8.9477 | -92.323
8.948 | -92.3238
8.9484 | -92.3251
8.9487 | -92.327
8.949 | -92.3282
8.9494 | -92.3286
8.9497 | -92.3294
8.9501 | -92.3297
8.9504 | -92.3298
8.9508 | -92.3306
8.9511 | -92.3307
8.9514 | -92.3309
8.9518 | -92.3317
8.9521 | -92.3323
8.9525 | -92.3332
8.9528 | -92.3344
8.9531 | -92.3343
8.9535 | -92.3341
8.9538 | -92.3347
8.9542 | -92.3361
8.9545 | -92.3368
8.9548 | -92.3374
8.9552 | -92.3383
8.9555 | -92.339
8.9559 | -92.3393
8.9562 | -92.3398
8.9565 | -92.3403
8.9569 | -92.3407
8.9572 | -92.3414
8.9576 | -92.3422
8.9579 | -92.3433
8.9582 | -92.3442
8.9586 | -92.3451
8.9589 | -92.3463
8.9593 | -92.3469
8.9596 | -92.3485
8.9599 | -92.3499
8.9603 | -92.3505
8.9606 | -92.351
8.961 | -92.3518
8.9613 | -92.3529
8.9616 | -92.3542
8.962 | -92.3553
8.9623 | -92.3559
8.9627 | -92.3568
8.963 | -92.3575
8.9634 | -92.3579
8.9637 | -92.3585
8.964 | -92.3594
8.9644 | -92.3605
8.9647 | -92.3616
8.9651 | -92.3626
8.9654 | -92.3647
8.9657 | -92.3661
8.9661 | -92.3669
8.9664 | -92.3677
8.9668 | -92.368
8.9671 | -92.3687
8.9674 | -92.3694
8.9678 | -92.3696
8.9681 | -92.3696
8.9685 | -92.3695
8.9688 | -92.3694
8.9691 | -92.3695
8.9695 | -92.3705
8.9698 | -92.3716
8.9702 | -92.3723
8.9705 | -92.3729
8.9708 | -92.3736
8.9712 | -92.3745
8.9715 | -92.3749
8.9719 | -92.3758
8.9722 | -92.3765
8.9725 | -92.3771
8.9729 | -92.3776
8.9732 | -92.3784
8.9736 | -92.3794
8.9739 | -92.3796
8.9742 | -92.3801
8.9746 | -92.3811
8.9749 | -92.3825
8.9753 | -92.3835
8.9756 | -92.3847
8.9759 | -92.3853
8.9763 | -92.386
8.9766 | -92.3867
8.977 | -92.3871
8.9773 | -92.388
8.9777 | -92.3886
8.978 | -92.3893
8.9783 | -92.3903
8.9787 | -92.3921
8.979 | -92.3928
8.9794 | -92.3935
8.9797 | -92.3939
8.98 | -92.394
8.9804 | -92.3943
8.9807 | -92.3944
8.9811 | -92.3945
8.9814 | -92.395
8.9817 | -92.3955
8.9821 | -92.3961
8.9824 | -92.3968
8.9828 | -92.3972
8.9831 | -92.3982
8.9834 | -92.3988
8.9838 | -92.3992
8.9841 | -92.3998
8.9845 | -92.4005
8.9848 | -92.4011
8.9851 | -92.4018
8.9855 | -92.4025
8.9858 | -92.4028
8.9862 | -92.4037
8.9865 | -92.4045
8.9868 | -92.404
8.9872 | -92.4043
8.9875 | -92.405
8.9879 | -92.4055
8.9882 | -92.4055
8.9885 | -92.4054
8.9889 | -92.4063
8.9892 | -92.4083
8.9896 | -92.4103
8.9899 | -92.4114
8.9902 | -92.412
8.9906 | -92.4126
8.9909 | -92.4131
8.9913 | -92.4136
8.9916 | -92.4141
8.992 | -92.4146
8.9923 | -92.4149
8.9926 | -92.4152
8.993 | -92.416
8.9933 | -92.4166
8.9937 | -92.4176
8.994 | -92.4183
8.9943 | -92.4187
8.9947 | -92.4191
8.995 | -92.4194
8.9954 | -92.42
8.9957 | -92.4204
8.996 | -92.4209
8.9964 | -92.4213
8.9967 | -92.4217
8.9971 | -92.4231
8.9974 | -92.425
8.9977 | -92.4269
8.9981 | -92.4278
8.9984 | -92.4282
8.9988 | -92.4287
8.9991 | -92.4301
8.9994 | -92.4309
8.9998 | -92.4314
9.0001 | -92.4321
9.0005 | -92.4329
9.0008 | -92.4332
9.0011 | -92.4332
9.0015 | -92.4327
9.0018 | -92.4335
9.0022 | -92.4345
9.0025 | -92.435
9.0028 | -92.4354
9.0032 | -92.436
9.0035 | -92.4371
9.0039 | -92.4386
9.0042 | -92.4399
9.0045 | -92.4399
9.0049 | -92.4392
9.0052 | -92.4391
9.0056 | -92.4395
9.0059 | -92.4402
9.0063 | -92.4406
9.0066 | -92.441
9.0069 | -92.4413
9.0073 | -92.4425
9.0076 | -92.4434
9.008 | -92.4439
9.0083 | -92.4435
9.0086 | -92.4439
9.009 | -92.444
9.0093 | -92.4444
9.0097 | -92.4456
9.01 | -92.4462
9.0103 | -92.4467
9.0107 | -92.4472
9.011 | -92.4479
9.0114 | -92.4482
9.0117 | -92.4485
9.012 | -92.4487
9.0124 | -92.449
9.0127 | -92.4494
9.0131 | -92.45
9.0134 | -92.4504
9.0137 | -92.4516
9.0141 | -92.4525
9.0144 | -92.4521
9.0148 | -92.4518
9.0151 | -92.4519
9.0154 | -92.4527
9.0158 | -92.4536
9.0161 | -92.4545
9.0165 | -92.4552
9.0168 | -92.4556
9.0171 | -92.4563
9.0175 | -92.4569
9.0178 | -92.4575
9.0182 | -92.4572
9.0185 | -92.4578
9.0188 | -92.458
9.0192 | -92.4588
9.0195 | -92.4592
9.0199 | -92.4595
9.0202 | -92.4597
9.0206 | -92.4599
9.0209 | -92.4602
9.0212 | -92.4607
9.0216 | -92.4611
9.0219 | -92.4614
9.0223 | -92.4613
9.0226 | -92.461
9.0229 | -92.4613
9.0233 | -92.4621
9.0236 | -92.4627
9.024 | -92.4632
9.0243 | -92.4638
9.0246 | -92.4644
9.025 | -92.4651
9.0253 | -92.4657
9.0257 | -92.4662
9.026 | -92.4668
9.0263 | -92.4671
9.0267 | -92.4677
9.027 | -92.468
9.0274 | -92.4682
9.0277 | -92.4686
9.028 | -92.4692
9.0284 | -92.4702
9.0287 | -92.4714
9.0291 | -92.472
9.0294 | -92.473
9.0297 | -92.4739
9.0301 | -92.4747
9.0304 | -92.4751
9.0308 | -92.4753
9.0311 | -92.4762
9.0314 | -92.4766
9.0318 | -92.4769
9.0321 | -92.4769
9.0325 | -92.4773
9.0328 | -92.4779
9.0331 | -92.4781
9.0335 | -92.4782
9.0338 | -92.4787
9.0342 | -92.4792
9.0345 | -92.4795
9.0349 | -92.48
9.0352 | -92.4807
9.0355 | -92.4812
9.0359 | -92.4819
9.0362 | -92.4829
9.0366 | -92.4835
9.0369 | -92.4838
9.0372 | -92.4841
9.0376 | -92.4844
9.0379 | -92.4849
9.0383 | -92.4856
9.0386 | -92.4871
9.0389 | -92.4875
9.0393 | -92.4878
9.0396 | -92.4882
9.04 | -92.4894
9.0403 | -92.4905
9.0406 | -92.4906
9.041 | -92.4902
9.0413 | -92.4907
9.0417 | -92.4914
9.042 | -92.4917
9.0423 | -92.492
9.0427 | -92.4927
9.043 | -92.4934
9.0434 | -92.4943
9.0437 | -92.4948
9.044 | -92.4948
9.0444 | -92.4949
9.0447 | -92.4949
9.0451 | -92.495
9.0454 | -92.4954
9.0457 | -92.4957
9.0461 | -92.4958
9.0464 | -92.4955
9.0468 | -92.4959
9.0471 | -92.4962
9.0474 | -92.4964
9.0478 | -92.4962
9.0481 | -92.4965
9.0485 | -92.4974
9.0488 | -92.4981
9.0492 | -92.4986
9.0495 | -92.4994
9.0498 | -92.5001
9.0502 | -92.5008
9.0505 | -92.5017
9.0509 | -92.5021
9.0512 | -92.5025
9.0515 | -92.5029
9.0519 | -92.5029
9.0522 | -92.5031
9.0526 | -92.5033
9.0529 | -92.5034
9.0532 | -92.5042
9.0536 | -92.5048
9.0539 | -92.5052
9.0543 | -92.5057
9.0546 | -92.5059
9.0549 | -92.5062
9.0553 | -92.5063
9.0556 | -92.5067
9.056 | -92.5075
9.0563 | -92.5079
9.0566 | -92.508
9.057 | -92.5085
9.0573 | -92.5086
9.0577 | -92.5092
9.058 | -92.5098
9.0583 | -92.5101
9.0587 | -92.5103
9.059 | -92.5104
9.0594 | -92.5108
9.0597 | -92.5117
9.06 | -92.5123
9.0604 | -92.5128
9.0607 | -92.5128
9.0611 | -92.5129
9.0614 | -92.5125
9.0618 | -92.512
9.0621 | -92.5121
9.0624 | -92.5121
9.0628 | -92.5118
9.0631 | -92.5121
9.0635 | -92.5121
9.0638 | -92.5128
9.0641 | -92.5146
9.0645 | -92.5155
9.0648 | -92.5151
9.0652 | -92.5152
9.0655 | -92.5148
9.0658 | -92.5136
9.0662 | -92.5126
9.0665 | -92.5131
9.0669 | -92.5139
9.0672 | -92.5144
9.0675 | -92.5152
9.0679 | -92.5154
9.0682 | -92.5161
9.0686 | -92.5171
9.0689 | -92.5177
9.0692 | -92.5184
9.0696 | -92.5195
9.0699 | -92.5201
9.0703 | -92.5201
9.0706 | -92.5206
9.0709 | -92.5208
9.0713 | -92.5202
9.0716 | -92.5203
9.072 | -92.5202
9.0723 | -92.5202
9.0726 | -92.5206
9.073 | -92.5207
9.0733 | -92.5208
9.0737 | -92.5211
9.074 | -92.5211
9.0743 | -92.5214
9.0747 | -92.5214
9.075 | -92.5214
9.0754 | -92.5216
9.0757 | -92.5214
9.0761 | -92.521
9.0764 | -92.5211
9.0767 | -92.5209
9.0771 | -92.5201
9.0774 | -92.5195
9.0778 | -92.519
9.0781 | -92.5191
9.0784 | -92.5194
9.0788 | -92.5193
9.0791 | -92.5196
9.0795 | -92.5203
9.0798 | -92.5205
9.0801 | -92.5208
9.0805 | -92.522
9.0808 | -92.5229
9.0812 | -92.5234
9.0815 | -92.5239
9.0818 | -92.5239
9.0822 | -92.5241
9.0825 | -92.5247
9.0829 | -92.525
9.0832 | -92.5251
9.0835 | -92.5252
9.0839 | -92.5254
9.0842 | -92.526
9.0846 | -92.5268
9.0849 | -92.527
9.0852 | -92.5269
9.0856 | -92.5269
9.0859 | -92.5268
9.0863 | -92.5271
9.0866 | -92.5275
9.0869 | -92.528
9.0873 | -92.5287
9.0876 | -92.5304
9.088 | -92.532
9.0883 | -92.5319
9.0886 | -92.5319
9.089 | -92.5323
9.0893 | -92.5323
9.0897 | -92.5332
9.09 | -92.5341
9.0904 | -92.5345
9.0907 | -92.535
9.091 | -92.5357
9.0914 | -92.5359
9.0917 | -92.5362
9.0921 | -92.537
9.0924 | -92.5369
9.0927 | -92.5364
9.0931 | -92.5357
9.0934 | -92.5348
9.0938 | -92.5343
9.0941 | -92.5347
9.0944 | -92.5349
9.0948 | -92.5349
9.0951 | -92.5347
9.0955 | -92.535
9.0958 | -92.5352
9.0961 | -92.5353
9.0965 | -92.5355
9.0968 | -92.5361
9.0972 | -92.5361
9.0975 | -92.5363
9.0978 | -92.5368
9.0982 | -92.5371
9.0985 | -92.5379
9.0989 | -92.5383
9.0992 | -92.5384
9.0995 | -92.5387
9.0999 | -92.5392
9.1002 | -92.5394
9.1006 | -92.5398
9.1009 | -92.5405
9.1012 | -92.5403
9.1016 | -92.5397
9.1019 | -92.5406
9.1023 | -92.5413
9.1026 | -92.5415
9.1029 | -92.5417
9.1033 | -92.5422
9.1036 | -92.5425
9.104 | -92.5426
9.1043 | -92.5426
9.1047 | -92.5429
9.105 | -92.5438
9.1053 | -92.5443
9.1057 | -92.5444
9.106 | -92.5442
9.1064 | -92.5444
9.1067 | -92.5448
9.107 | -92.5449
9.1074 | -92.545
9.1077 | -92.5451
9.1081 | -92.5455
9.1084 | -92.5464
9.1087 | -92.5467
9.1091 | -92.5473
9.1094 | -92.5481
9.1098 | -92.5483
9.1101 | -92.549
9.1104 | -92.5504
9.1108 | -92.5506
9.1111 | -92.5503
9.1115 | -92.5505
9.1118 | -92.5507
9.1121 | -92.5512
9.1125 | -92.5516
9.1128 | -92.5514
9.1132 | -92.5512
9.1135 | -92.5515
9.1138 | -92.5516
9.1142 | -92.5514
9.1145 | -92.5516
9.1149 | -92.5519
9.1152 | -92.5521
9.1155 | -92.5524
9.1159 | -92.5525
9.1162 | -92.5532
9.1166 | -92.5538
9.1169 | -92.5541
9.1173 | -92.5549
9.1176 | -92.5556
9.1179 | -92.5566
9.1183 | -92.5583
9.1186 | -92.5586
9.119 | -92.5588
9.1193 | -92.5592
9.1196 | -92.5602
9.12 | -92.5607
9.1203 | -92.5608
9.1207 | -92.5609
9.121 | -92.5609
9.1213 | -92.5612
9.1217 | -92.5612
9.122 | -92.5614
9.1224 | -92.5617
9.1227 | -92.5618
9.123 | -92.5621
9.1234 | -92.5623
9.1237 | -92.562
9.1241 | -92.5608
9.1244 | -92.5601
9.1247 | -92.5603
9.1251 | -92.5597
9.1254 | -92.5594
9.1258 | -92.5604
9.1261 | -92.561
9.1264 | -92.5616
9.1268 | -92.5622
9.1271 | -92.5624
9.1275 | -92.5634
9.1278 | -92.564
9.1281 | -92.5643
9.1285 | -92.565
9.1288 | -92.565
9.1292 | -92.5649
9.1295 | -92.5645
9.1298 | -92.5644
9.1302 | -92.5642
9.1305 | -92.5636
9.1309 | -92.5632
9.1312 | -92.5632
9.1316 | -92.5633
9.1319 | -92.5639
9.1322 | -92.5643
9.1326 | -92.5645
9.1329 | -92.5647
9.1333 | -92.5648
9.1336 | -92.5654
9.1339 | -92.5656
9.1343 | -92.5657
9.1346 | -92.566
9.135 | -92.5657
9.1353 | -92.5654
9.1356 | -92.5658
9.136 | -92.5658
9.1363 | -92.5657
9.1367 | -92.5659
9.137 | -92.567
9.1373 | -92.5674
9.1377 | -92.5676
9.138 | -92.5679
9.1384 | -92.5685
9.1387 | -92.5693
9.139 | -92.5694
9.1394 | -92.5696
9.1397 | -92.5698
9.1401 | -92.5701
9.1404 | -92.5702
9.1407 | -92.5698
9.1411 | -92.5701
9.1414 | -92.5707
9.1418 | -92.5707
9.1421 | -92.5709
9.1424 | -92.571
9.1428 | -92.5707
9.1431 | -92.57
9.1435 | -92.5688
9.1438 | -92.5683
9.1441 | -92.5685
9.1445 | -92.5691
9.1448 | -92.5701
9.1452 | -92.5703
9.1455 | -92.5693
9.1459 | -92.5685
9.1462 | -92.5687
9.1465 | -92.5685
9.1469 | -92.5685
9.1472 | -92.57
9.1476 | -92.5714
9.1479 | -92.5719
9.1482 | -92.5721
9.1486 | -92.5724
9.1489 | -92.5727
9.1493 | -92.5732
9.1496 | -92.5734
9.1499 | -92.5731
9.1503 | -92.5733
9.1506 | -92.5737
9.151 | -92.574
9.1513 | -92.5744
9.1516 | -92.5751
9.152 | -92.5754
9.1523 | -92.5755
9.1527 | -92.5756
9.153 | -92.5758
9.1533 | -92.5757
9.1537 | -92.576
9.154 | -92.576
9.1544 | -92.5767
9.1547 | -92.5772
9.155 | -92.5772
9.1554 | -92.5773
9.1557 | -92.5778
9.1561 | -92.5782
9.1564 | -92.5785
9.1567 | -92.579
9.1571 | -92.5792
9.1574 | -92.5796
9.1578 | -92.5799
9.1581 | -92.5802
9.1585 | -92.5798
9.1588 | -92.5789
9.1591 | -92.5789
9.1595 | -92.5794
9.1598 | -92.5797
9.1602 | -92.5797
9.1605 | -92.5798
9.1608 | -92.5799
9.1612 | -92.5801
9.1615 | -92.5804
9.1619 | -92.5807
9.1622 | -92.5808
9.1625 | -92.581
9.1629 | -92.5811
9.1632 | -92.5814
9.1636 | -92.5817
9.1639 | -92.5823
9.1642 | -92.5823
9.1646 | -92.5823
9.1649 | -92.5827
9.1653 | -92.5827
9.1656 | -92.5822
9.1659 | -92.5826
9.1663 | -92.5828
9.1666 | -92.5832
9.167 | -92.5838
9.1673 | -92.5834
9.1676 | -92.5828
9.168 | -92.5828
9.1683 | -92.5831
9.1687 | -92.5838
9.169 | -92.5837
9.1693 | -92.5836
9.1697 | -92.5842
9.17 | -92.5847
9.1704 | -92.5845
9.1707 | -92.5845
9.171 | -92.585
9.1714 | -92.5852
9.1717 | -92.5854
9.1721 | -92.5857
9.1724 | -92.586
9.1728 | -92.5863
9.1731 | -92.5865
9.1734 | -92.5874
9.1738 | -92.5884
9.1741 | -92.5886
9.1745 | -92.5889
9.1748 | -92.5893
9.1751 | -92.5901
9.1755 | -92.5906
9.1758 | -92.591
9.1762 | -92.5913
9.1765 | -92.5917
9.1768 | -92.5921
9.1772 | -92.5931
9.1775 | -92.5945
9.1779 | -92.5949
9.1782 | -92.5947
9.1785 | -92.5948
9.1789 | -92.5954
9.1792 | -92.5955
9.1796 | -92.5955
9.1799 | -92.5958
9.1802 | -92.5961
9.1806 | -92.5964
9.1809 | -92.5972
9.1813 | -92.5978
9.1816 | -92.5979
9.1819 | -92.598
9.1823 | -92.5985
9.1826 | -92.5987
9.183 | -92.5986
9.1833 | -92.5988
9.1836 | -92.5988
9.184 | -92.5987
9.1843 | -92.5989
9.1847 | -92.5992
9.185 | -92.5998
9.1853 | -92.6005
9.1857 | -92.6008
9.186 | -92.6011
9.1864 | -92.6013
9.1867 | -92.6017
9.1871 | -92.602
9.1874 | -92.6024
9.1877 | -92.6028
9.1881 | -92.6029
9.1884 | -92.6031
9.1888 | -92.6033
9.1891 | -92.6037
9.1894 | -92.6041
9.1898 | -92.6045
9.1901 | -92.605
9.1905 | -92.606
9.1908 | -92.6059
9.1911 | -92.6064
9.1915 | -92.6068
9.1918 | -92.6068
9.1922 | -92.6072
9.1925 | -92.6076
9.1928 | -92.6078
9.1932 | -92.6084
9.1935 | -92.6093
9.1939 | -92.6101
9.1942 | -92.6109
9.1945 | -92.6114
9.1949 | -92.6121
9.1952 | -92.6124
9.1956 | -92.6131
9.1959 | -92.6136
9.1962 | -92.6141
9.1966 | -92.6144
9.1969 | -92.6145
9.1973 | -92.6147
9.1976 | -92.615
9.1979 | -92.6155
9.1983 | -92.6159
9.1986 | -92.6167
9.199 | -92.6172
9.1993 | -92.6173
9.1996 | -92.6175
9.2 | -92.618
9.2003 | -92.6185
9.2007 | -92.6188
9.201 | -92.6197
9.2014 | -92.6213
9.2017 | -92.6226
9.202 | -92.6235
9.2024 | -92.6239
9.2027 | -92.6248
9.2031 | -92.6253
9.2034 | -92.6255
9.2037 | -92.6257
9.2041 | -92.6257
9.2044 | -92.6262
9.2048 | -92.627
9.2051 | -92.6278
9.2054 | -92.6283
9.2058 | -92.6285
9.2061 | -92.6293
9.2065 | -92.6305
9.2068 | -92.6311
9.2071 | -92.6317
9.2075 | -92.6325
9.2078 | -92.6327
9.2082 | -92.633
9.2085 | -92.633
9.2088 | -92.6332
9.2092 | -92.6331
9.2095 | -92.6331
9.2099 | -92.6334
9.2102 | -92.6338
9.2105 | -92.6342
9.2109 | -92.6345
9.2112 | -92.6348
9.2116 | -92.6355
9.2119 | -92.6357
9.2122 | -92.6355
9.2126 | -92.6359
9.2129 | -92.6365
9.2133 | -92.6369
9.2136 | -92.6374
9.2139 | -92.6381
9.2143 | -92.6391
9.2146 | -92.6399
9.215 | -92.6411
9.2153 | -92.642
9.2157 | -92.6423
9.216 | -92.6424
9.2163 | -92.6425
9.2167 | -92.6428
9.217 | -92.643
9.2174 | -92.6435
9.2177 | -92.6441
9.218 | -92.6445
9.2184 | -92.6451
9.2187 | -92.6456
9.2191 | -92.6458
9.2194 | -92.6457
9.2197 | -92.6461
9.2201 | -92.6469
9.2204 | -92.648
9.2208 | -92.6498
9.2211 | -92.6508
9.2214 | -92.651
9.2218 | -92.6512
9.2221 | -92.6514
9.2225 | -92.6516
9.2228 | -92.6515
9.2231 | -92.6516
9.2235 | -92.6521
9.2238 | -92.6525
9.2242 | -92.6533
9.2245 | -92.6543
9.2248 | -92.6556
9.2252 | -92.6562
9.2255 | -92.6567
9.2259 | -92.6575
9.2262 | -92.658
9.2265 | -92.6583
9.2269 | -92.6587
9.2272 | -92.6591
9.2276 | -92.6598
9.2279 | -92.6603
9.2282 | -92.6608
9.2286 | -92.6618
9.2289 | -92.6626
9.2293 | -92.6631
9.2296 | -92.6641
9.23 | -92.6647
9.2303 | -92.6649
9.2306 | -92.6654
9.231 | -92.6664
9.2313 | -92.6672
9.2317 | -92.6676
9.232 | -92.6677
9.2323 | -92.6687
9.2327 | -92.6695
9.233 | -92.6698
9.2334 | -92.6703
9.2337 | -92.6709
9.234 | -92.6712
9.2344 | -92.672
9.2347 | -92.6725
9.2351 | -92.673
9.2354 | -92.6735
9.2357 | -92.6742
9.2361 | -92.6748
9.2364 | -92.6753
9.2368 | -92.6764
9.2371 | -92.6772
9.2374 | -92.6774
9.2378 | -92.6773
9.2381 | -92.6782
9.2385 | -92.6801
9.2388 | -92.6807
9.2391 | -92.6812
9.2395 | -92.6818
9.2398 | -92.6823
9.2402 | -92.683
9.2405 | -92.6833
9.2408 | -92.683
9.2412 | -92.6829
9.2415 | -92.6833
9.2419 | -92.6835
9.2422 | -92.684
9.2425 | -92.6855
9.2429 | -92.687
9.2432 | -92.6875
9.2436 | -92.6878
9.2439 | -92.688
9.2443 | -92.6881
9.2446 | -92.6882
9.2449 | -92.6885
9.2453 | -92.6887
9.2456 | -92.6892
9.246 | -92.6901
9.2463 | -92.6906
9.2466 | -92.691
9.247 | -92.6915
9.2473 | -92.692
9.2477 | -92.6925
9.248 | -92.6936
9.2483 | -92.6949
9.2487 | -92.6955
9.249 | -92.6963
9.2494 | -92.6975
9.2497 | -92.6984
9.25 | -92.6988
9.2504 | -92.6993
9.2507 | -92.7001
9.2511 | -92.7009
9.2514 | -92.7012
9.2517 | -92.7017
9.2521 | -92.7024
9.2524 | -92.7035
9.2528 | -92.7044
9.2531 | -92.705
9.2534 | -92.7052
9.2538 | -92.7066
9.2541 | -92.707
9.2545 | -92.7079
9.2548 | -92.7084
9.2551 | -92.709
9.2555 | -92.7095
9.2558 | -92.71
9.2562 | -92.7106
9.2565 | -92.7114
9.2569 | -92.712
9.2572 | -92.7126
9.2575 | -92.7131
9.2579 | -92.7129
9.2582 | -92.7121
9.2586 | -92.7119
9.2589 | -92.7124
9.2592 | -92.7128
9.2596 | -92.7132
9.2599 | -92.7136
9.2603 | -92.7141
9.2606 | -92.7147
9.2609 | -92.7152
9.2613 | -92.7159
9.2616 | -92.7157
9.262 | -92.7156
9.2623 | -92.7153
9.2626 | -92.7157
9.263 | -92.7163
9.2633 | -92.7172
9.2637 | -92.718
9.264 | -92.7192
9.2643 | -92.7201
9.2647 | -92.7205
9.265 | -92.7211
9.2654 | -92.722
9.2657 | -92.7226
9.266 | -92.7229
9.2664 | -92.7235
9.2667 | -92.7243
9.2671 | -92.7248
9.2674 | -92.7255
9.2677 | -92.726
9.2681 | -92.7265
9.2684 | -92.7278
9.2688 | -92.7294
9.2691 | -92.7307
9.2694 | -92.7319
9.2698 | -92.7324
9.2701 | -92.7329
9.2705 | -92.7331
9.2708 | -92.7337
9.2712 | -92.735
9.2715 | -92.737
9.2718 | -92.7378
9.2722 | -92.7381
9.2725 | -92.7382
9.2729 | -92.7387
9.2732 | -92.7392
9.2735 | -92.7403
9.2739 | -92.7405
9.2742 | -92.7407
9.2746 | -92.7411
9.2749 | -92.7419
9.2752 | -92.7424
9.2756 | -92.7426
9.2759 | -92.7432
9.2763 | -92.744
9.2766 | -92.745
9.2769 | -92.7456
9.2773 | -92.7462
9.2776 | -92.7466
9.278 | -92.7474
9.2783 | -92.7485
9.2786 | -92.7487
9.279 | -92.7485
9.2793 | -92.749
9.2797 | -92.7493
9.28 | -92.7492
9.2803 | -92.7489
9.2807 | -92.7496
9.281 | -92.7498
9.2814 | -92.7471
9.2817 | -92.7451
9.282 | -92.7457
9.2824 | -92.7464
9.2827 | -92.747
9.2831 | -92.7477
9.2834 | -92.7487
9.2837 | -92.7493
9.2841 | -92.75
9.2844 | -92.7504
9.2848 | -92.7511
9.2851 | -92.7517
9.2855 | -92.7523
9.2858 | -92.7527
9.2861 | -92.7524
9.2865 | -92.7526
9.2868 | -92.7532
9.2872 | -92.7534
9.2875 | -92.7541
9.2878 | -92.7547
9.2882 | -92.755
9.2885 | -92.7554
9.2889 | -92.7563
9.2892 | -92.7572
9.2895 | -92.7578
9.2899 | -92.7585
9.2902 | -92.7592
9.2906 | -92.7598
9.2909 | -92.7605
9.2912 | -92.761
9.2916 | -92.7614
9.2919 | -92.762
9.2923 | -92.7627
9.2926 | -92.7633
9.2929 | -92.764
9.2933 | -92.7647
9.2936 | -92.7653
9.294 | -92.7664
9.2943 | -92.7672
9.2946 | -92.768
9.295 | -92.7686
9.2953 | -92.7698
9.2957 | -92.7708
9.296 | -92.7716
9.2963 | -92.7729
9.2967 | -92.774
9.297 | -92.7747
9.2974 | -92.7754
9.2977 | -92.776
9.298 | -92.7766
9.2984 | -92.7771
9.2987 | -92.7774
9.2991 | -92.7777
9.2994 | -92.7783
9.2998 | -92.7791
9.3001 | -92.7803
9.3004 | -92.7809
9.3008 | -92.7813
9.3011 | -92.7819
9.3015 | -92.7826
9.3018 | -92.7835
9.3021 | -92.784
9.3025 | -92.7841
9.3028 | -92.7841
9.3032 | -92.7851
9.3035 | -92.7857
9.3038 | -92.7864
9.3042 | -92.7871
9.3045 | -92.7879
9.3049 | -92.7884
9.3052 | -92.7891
9.3055 | -92.7898
9.3059 | -92.7903
9.3062 | -92.7909
9.3066 | -92.7915
9.3069 | -92.7922
9.3072 | -92.793
9.3076 | -92.7939
9.3079 | -92.7946
9.3083 | -92.7954
9.3086 | -92.7961
9.3089 | -92.7968
9.3093 | -92.7975
9.3096 | -92.7986
9.31 | -92.7997
9.3103 | -92.8007
9.3106 | -92.8013
9.311 | -92.8023
9.3113 | -92.8035
9.3117 | -92.8046
9.312 | -92.8053
9.3124 | -92.8061
9.3127 | -92.8069
9.313 | -92.8072
9.3134 | -92.8078
9.3137 | -92.8079
9.3141 | -92.808
9.3144 | -92.8085
9.3147 | -92.8091
9.3151 | -92.809
9.3154 | -92.8093
9.3158 | -92.8101
9.3161 | -92.8115
9.3164 | -92.8125
9.3168 | -92.8135
9.3171 | -92.8146
9.3175 | -92.8149
9.3178 | -92.8149
9.3181 | -92.8154
9.3185 | -92.8164
9.3188 | -92.8169
9.3192 | -92.8171
9.3195 | -92.8178
9.3198 | -92.8185
9.3202 | -92.8188
9.3205 | -92.8192
9.3209 | -92.8198
9.3212 | -92.8204
9.3215 | -92.8215
9.3219 | -92.8229
9.3222 | -92.8234
9.3226 | -92.8238
9.3229 | -92.8242
9.3232 | -92.8246
9.3236 | -92.8253
9.3239 | -92.8261
9.3243 | -92.8267
9.3246 | -92.8276
9.3249 | -92.8283
9.3253 | -92.8285
9.3256 | -92.829
9.326 | -92.8296
9.3263 | -92.8303
9.3267 | -92.831
9.327 | -92.8315
9.3273 | -92.8321
9.3277 | -92.8326
9.328 | -92.8331
9.3284 | -92.8341
9.3287 | -92.8353
9.329 | -92.8359
9.3294 | -92.8368
9.3297 | -92.8372
9.3301 | -92.8384
9.3304 | -92.8392
9.3307 | -92.8399
9.3311 | -92.8406
9.3314 | -92.841
9.3318 | -92.8417
9.3321 | -92.8423
9.3324 | -92.8431
9.3328 | -92.8438
9.3331 | -92.8444
9.3335 | -92.8448
9.3338 | -92.8449
9.3341 | -92.8453
9.3345 | -92.8458
9.3348 | -92.8464
9.3352 | -92.847
9.3355 | -92.8476
9.3358 | -92.848
9.3362 | -92.8483
9.3365 | -92.848
9.3369 | -92.848
9.3372 | -92.8484
9.3375 | -92.8485
9.3379 | -92.848
9.3382 | -92.8483
9.3386 | -92.8495
9.3389 | -92.8502
9.3392 | -92.8507
9.3396 | -92.8512
9.3399 | -92.852
9.3403 | -92.8527
9.3406 | -92.853
9.341 | -92.8529
9.3413 | -92.853
9.3416 | -92.8535
9.342 | -92.854
9.3423 | -92.8549
9.3427 | -92.8557
9.343 | -92.8566
9.3433 | -92.8571
9.3437 | -92.857
9.344 | -92.8566
9.3444 | -92.857
9.3447 | -92.8574
9.345 | -92.8582
9.3454 | -92.8591
9.3457 | -92.8591
9.3461 | -92.8587
9.3464 | -92.8591
9.3467 | -92.8597
9.3471 | -92.8599
9.3474 | -92.8607
9.3478 | -92.8619
9.3481 | -92.8631
9.3484 | -92.864
9.3488 | -92.8645
9.3491 | -92.8638
9.3495 | -92.8644
9.3498 | -92.8655
9.3501 | -92.8663
9.3505 | -92.8667
9.3508 | -92.8673
9.3512 | -92.8677
9.3515 | -92.8684
9.3518 | -92.8691
9.3522 | -92.87
9.3525 | -92.8701
9.3529 | -92.8701
9.3532 | -92.8705
9.3535 | -92.871
9.3539 | -92.8716
9.3542 | -92.8721
9.3546 | -92.8729
9.3549 | -92.8737
9.3553 | -92.8741
9.3556 | -92.8745
9.3559 | -92.8756
9.3563 | -92.8771
9.3566 | -92.8781
9.357 | -92.8794
9.3573 | -92.8806
9.3576 | -92.8816
9.358 | -92.8826
9.3583 | -92.8834
9.3587 | -92.8838
9.359 | -92.8844
9.3593 | -92.8849
9.3597 | -92.8858
9.36 | -92.8866
9.3604 | -92.8874
9.3607 | -92.8879
9.361 | -92.8889
9.3614 | -92.8901
9.3617 | -92.8901
9.3621 | -92.8901
9.3624 | -92.8903
9.3627 | -92.8908
9.3631 | -92.8916
9.3634 | -92.8924
9.3638 | -92.8935
9.3641 | -92.8948
9.3644 | -92.8963
9.3648 | -92.8965
9.3651 | -92.8966
9.3655 | -92.8974
9.3658 | -92.898
9.3661 | -92.8985
9.3665 | -92.899
9.3668 | -92.8998
9.3672 | -92.9004
9.3675 | -92.9007
9.3679 | -92.9009
9.3682 | -92.9006
9.3685 | -92.901
9.3689 | -92.9017
9.3692 | -92.9022
9.3696 | -92.9025
9.3699 | -92.9029
9.3702 | -92.9033
9.3706 | -92.9033
9.3709 | -92.9031
9.3713 | -92.9038
9.3716 | -92.9047
9.3719 | -92.9051
9.3723 | -92.9054
9.3726 | -92.9061
9.373 | -92.9067
9.3733 | -92.9078
9.3736 | -92.9089
9.374 | -92.9094
9.3743 | -92.91
9.3747 | -92.9108
9.375 | -92.9123
9.3753 | -92.9131
9.3757 | -92.9136
9.376 | -92.9143
9.3764 | -92.9151
9.3767 | -92.916
9.377 | -92.917
9.3774 | -92.9181
9.3777 | -92.9187
9.3781 | -92.9194
9.3784 | -92.9199
9.3787 | -92.9206
9.3791 | -92.9216
9.3794 | -92.9224
9.3798 | -92.9227
9.3801 | -92.923
9.3804 | -92.9236
9.3808 | -92.924
9.3811 | -92.925
9.3815 | -92.9254
9.3818 | -92.9262
9.3821 | -92.9268
9.3825 | -92.9273
9.3828 | -92.9277
9.3832 | -92.9283
9.3835 | -92.9294
9.3839 | -92.931
9.3842 | -92.9318
9.3845 | -92.9326
9.3849 | -92.9331
9.3852 | -92.9337
9.3856 | -92.9345
9.3859 | -92.936
9.3862 | -92.9366
9.3866 | -92.9366
9.3869 | -92.9372
9.3873 | -92.9375
9.3876 | -92.9375
9.3879 | -92.9377
9.3883 | -92.938
9.3886 | -92.9382
9.389 | -92.9388
9.3893 | -92.9394
9.3896 | -92.9399
9.39 | -92.9406
9.3903 | -92.941
9.3907 | -92.9416
9.391 | -92.9424
9.3913 | -92.9424
9.3917 | -92.9431
9.392 | -92.9443
9.3924 | -92.9444
9.3927 | -92.9449
9.393 | -92.945
9.3934 | -92.9443
9.3937 | -92.9443
9.3941 | -92.9448
9.3944 | -92.9452
9.3947 | -92.9458
9.3951 | -92.9472
9.3954 | -92.9488
9.3958 | -92.95
9.3961 | -92.951
9.3965 | -92.9517
9.3968 | -92.9525
9.3971 | -92.9536
9.3975 | -92.9548
9.3978 | -92.9555
9.3982 | -92.956
9.3985 | -92.9563
9.3988 | -92.9566
9.3992 | -92.9571
9.3995 | -92.9575
9.3999 | -92.9576
9.4002 | -92.9579
9.4005 | -92.9584
9.4009 | -92.9592
9.4012 | -92.9592
9.4016 | -92.9595
9.4019 | -92.96
9.4022 | -92.9602
9.4026 | -92.9608
9.4029 | -92.962
9.4033 | -92.9632
9.4036 | -92.964
9.4039 | -92.9647
9.4043 | -92.965
9.4046 | -92.9653
9.405 | -92.9658
9.4053 | -92.9662
9.4056 | -92.9664
9.406 | -92.9666
9.4063 | -92.9669
9.4067 | -92.9673
9.407 | -92.968
9.4073 | -92.9686
9.4077 | -92.9695
9.408 | -92.9702
9.4084 | -92.9702
9.4087 | -92.9704
9.409 | -92.9707
9.4094 | -92.9715
9.4097 | -92.9723
9.4101 | -92.9732
9.4104 | -92.9736
9.4108 | -92.9738
9.4111 | -92.9743
9.4114 | -92.975
9.4118 | -92.9757
9.4121 | -92.976
9.4125 | -92.9763
9.4128 | -92.9769
9.4131 | -92.9775
9.4135 | -92.9779
9.4138 | -92.9785
9.4142 | -92.9788
9.4145 | -92.9789
9.4148 | -92.9794
9.4152 | -92.9805
9.4155 | -92.9817
9.4159 | -92.9829
9.4162 | -92.9832
9.4165 | -92.9839
9.4169 | -92.9842
9.4172 | -92.9839
9.4176 | -92.9843
9.4179 | -92.9846
9.4182 | -92.9849
9.4186 | -92.9856
9.4189 | -92.9864
9.4193 | -92.9868
9.4196 | -92.9877
9.4199 | -92.9885
9.4203 | -92.989
9.4206 | -92.9893
9.421 | -92.99
9.4213 | -92.9905
9.4216 | -92.991
9.422 | -92.9918
9.4223 | -92.9924
9.4227 | -92.993
9.423 | -92.9937
9.4233 | -92.9942
9.4237 | -92.9949
9.424 | -92.9952
9.4244 | -92.996
9.4247 | -92.9973
9.4251 | -92.9981
9.4254 | -92.9988
9.4257 | -92.9993
9.4261 | -93.0002
9.4264 | -93.0011
9.4268 | -93.0016
9.4271 | -93.0022
9.4274 | -93.0033
9.4278 | -93.004
9.4281 | -93.0045
9.4285 | -93.0047
9.4288 | -93.0051
9.4291 | -93.0056
9.4295 | -93.0066
9.4298 | -93.0081
9.4302 | -93.0089
9.4305 | -93.0102
9.4308 | -93.0109
9.4312 | -93.013
9.4315 | -93.0143
9.4319 | -93.0146
9.4322 | -93.0148
9.4325 | -93.0149
9.4329 | -93.0153
9.4332 | -93.0156
9.4336 | -93.0163
9.4339 | -93.0169
9.4342 | -93.0171
9.4346 | -93.0172
9.4349 | -93.0178
9.4353 | -93.0185
9.4356 | -93.0183
9.4359 | -93.0189
9.4363 | -93.0191
9.4366 | -93.0195
9.437 | -93.0197
9.4373 | -93.0197
9.4376 | -93.0202
9.438 | -93.0205
9.4383 | -93.0206
9.4387 | -93.0207
9.439 | -93.021
9.4394 | -93.0215
9.4397 | -93.0217
9.44 | -93.0221
9.4404 | -93.0232
9.4407 | -93.0231
9.4411 | -93.0232
9.4414 | -93.0232
9.4417 | -93.0236
9.4421 | -93.0242
9.4424 | -93.0249
9.4428 | -93.0253
9.4431 | -93.026
9.4434 | -93.0261
9.4438 | -93.026
9.4441 | -93.0266
9.4445 | -93.0272
9.4448 | -93.0273
9.4451 | -93.0279
9.4455 | -93.0288
9.4458 | -93.0292
9.4462 | -93.0298
9.4465 | -93.0303
9.4468 | -93.031
9.4472 | -93.0317
9.4475 | -93.0321
9.4479 | -93.0325
9.4482 | -93.0327
9.4485 | -93.0328
9.4489 | -93.0334
9.4492 | -93.0343
9.4496 | -93.0349
9.4499 | -93.035
9.4502 | -93.0354
9.4506 | -93.0359
9.4509 | -93.0365
9.4513 | -93.0369
9.4516 | -93.0373
9.452 | -93.0376
9.4523 | -93.0382
9.4526 | -93.0392
9.453 | -93.04
9.4533 | -93.0404
9.4537 | -93.0408
9.454 | -93.0413
9.4543 | -93.0421
9.4547 | -93.0428
9.455 | -93.0433
9.4554 | -93.0439
9.4557 | -93.0447
9.456 | -93.0455
9.4564 | -93.0454
9.4567 | -93.0459
9.4571 | -93.0467
9.4574 | -93.0471
9.4577 | -93.0473
9.4581 | -93.0477
9.4584 | -93.0481
9.4588 | -93.0485
9.4591 | -93.0488
9.4594 | -93.0492
9.4598 | -93.0495
9.4601 | -93.0498
9.4605 | -93.0503
9.4608 | -93.0506
9.4611 | -93.0507
9.4615 | -93.0506
9.4618 | -93.0514
9.4622 | -93.0523
9.4625 | -93.0527
9.4628 | -93.053
9.4632 | -93.0535
9.4635 | -93.0542
9.4639 | -93.055
9.4642 | -93.055
9.4645 | -93.0553
9.4649 | -93.0556
9.4652 | -93.0562
9.4656 | -93.0569
9.4659 | -93.0576
9.4663 | -93.0588
9.4666 | -93.0597
9.4669 | -93.0601
9.4673 | -93.0604
9.4676 | -93.062
9.468 | -93.0628
9.4683 | -93.0632
9.4686 | -93.0636
9.469 | -93.0639
9.4693 | -93.0649
9.4697 | -93.0657
9.47 | -93.0662
9.4703 | -93.0665
9.4707 | -93.0671
9.471 | -93.0677
9.4714 | -93.0678
9.4717 | -93.0679
9.472 | -93.0686
9.4724 | -93.0693
9.4727 | -93.0699
9.4731 | -93.0704
9.4734 | -93.0709
9.4737 | -93.0711
9.4741 | -93.0715
9.4744 | -93.0716
9.4748 | -93.0721
9.4751 | -93.0736
9.4754 | -93.0754
9.4758 | -93.0773
9.4761 | -93.0787
9.4765 | -93.0795
9.4768 | -93.0802
9.4771 | -93.0807
9.4775 | -93.081
9.4778 | -93.0813
9.4782 | -93.0817
9.4785 | -93.0824
9.4788 | -93.0829
9.4792 | -93.0833
9.4795 | -93.0838
9.4799 | -93.0846
9.4802 | -93.0849
9.4806 | -93.0856
9.4809 | -93.0864
9.4812 | -93.0868
9.4816 | -93.0869
9.4819 | -93.0867
9.4823 | -93.0867
9.4826 | -93.087
9.4829 | -93.0876
9.4833 | -93.088
9.4836 | -93.0884
9.484 | -93.0893
9.4843 | -93.0901
9.4846 | -93.091
9.485 | -93.0917
9.4853 | -93.092
9.4857 | -93.0925
9.486 | -93.0936
9.4863 | -93.094
9.4867 | -93.0933
9.487 | -93.0937
9.4874 | -93.0946
9.4877 | -93.0952
9.488 | -93.0956
9.4884 | -93.0959
9.4887 | -93.096
9.4891 | -93.0964
9.4894 | -93.0967
9.4897 | -93.0968
9.4901 | -93.0976
9.4904 | -93.0989
9.4908 | -93.0997
9.4911 | -93.1
9.4914 | -93.1006
9.4918 | -93.1011
9.4921 | -93.1017
9.4925 | -93.1021
9.4928 | -93.102
9.4931 | -93.1021
9.4935 | -93.1024
9.4938 | -93.1025
9.4942 | -93.103
9.4945 | -93.1032
9.4949 | -93.1039
9.4952 | -93.1054
9.4955 | -93.1072
9.4959 | -93.108
9.4962 | -93.1086
9.4966 | -93.1097
9.4969 | -93.1103
9.4972 | -93.1107
9.4976 | -93.111
9.4979 | -93.1121
9.4983 | -93.1127
9.4986 | -93.1126
9.4989 | -93.1129
9.4993 | -93.1131
9.4996 | -93.1123
9.5 | -93.1121
9.5003 | -93.1123
9.5006 | -93.1128
9.501 | -93.1142
9.5013 | -93.1156
9.5017 | -93.1166
9.502 | -93.1173
9.5023 | -93.1181
9.5027 | -93.1182
9.503 | -93.1183
9.5034 | -93.1189
9.5037 | -93.1196
9.504 | -93.1201
9.5044 | -93.1207
9.5047 | -93.1216
9.5051 | -93.1224
9.5054 | -93.1226
9.5057 | -93.1229
9.5061 | -93.1235
9.5064 | -93.1242
9.5068 | -93.1248
9.5071 | -93.1251
9.5075 | -93.1259
9.5078 | -93.1272
9.5081 | -93.1275
9.5085 | -93.1279
9.5088 | -93.1281
9.5092 | -93.1282
9.5095 | -93.1292
9.5098 | -93.13
9.5102 | -93.1302
9.5105 | -93.1307
9.5109 | -93.1313
9.5112 | -93.1318
9.5115 | -93.1322
9.5119 | -93.1325
9.5122 | -93.1327
9.5126 | -93.1322
9.5129 | -93.1319
9.5132 | -93.1328
9.5136 | -93.1339
9.5139 | -93.1345
9.5143 | -93.1351
9.5146 | -93.1352
9.5149 | -93.1359
9.5153 | -93.1372
9.5156 | -93.138
9.516 | -93.1381
9.5163 | -93.1386
9.5166 | -93.1391
9.517 | -93.14
9.5173 | -93.1409
9.5177 | -93.1411
9.518 | -93.1419
9.5183 | -93.1432
9.5187 | -93.1438
9.519 | -93.1447
9.5194 | -93.1452
9.5197 | -93.1456
9.52 | -93.1463
9.5204 | -93.1471
9.5207 | -93.1477
9.5211 | -93.1482
9.5214 | -93.1488
9.5218 | -93.1494
9.5221 | -93.1505
9.5224 | -93.1513
9.5228 | -93.1514
9.5231 | -93.1514
9.5235 | -93.1518
9.5238 | -93.1521
9.5241 | -93.1522
9.5245 | -93.1525
9.5248 | -93.1531
9.5252 | -93.1547
9.5255 | -93.1554
9.5258 | -93.1558
9.5262 | -93.1561
9.5265 | -93.1563
9.5269 | -93.157
9.5272 | -93.1576
9.5275 | -93.1574
9.5279 | -93.1577
9.5282 | -93.1584
9.5286 | -93.159
9.5289 | -93.16
9.5292 | -93.1603
9.5296 | -93.1606
9.5299 | -93.161
9.5303 | -93.1614
9.5306 | -93.1614
9.5309 | -93.1612
9.5313 | -93.1605
9.5316 | -93.1602
9.532 | -93.1607
9.5323 | -93.1606
9.5326 | -93.1616
9.533 | -93.1628
9.5333 | -93.1636
9.5337 | -93.1645
9.534 | -93.1653
9.5343 | -93.1659
9.5347 | -93.1665
9.535 | -93.1668
9.5354 | -93.1673
9.5357 | -93.1682
9.5361 | -93.1688
9.5364 | -93.1688
9.5367 | -93.169
9.5371 | -93.1693
9.5374 | -93.1694
9.5378 | -93.1698
9.5381 | -93.1701
9.5384 | -93.1702
9.5388 | -93.1709
9.5391 | -93.1713
9.5395 | -93.1717
9.5398 | -93.1722
9.5401 | -93.1732
9.5405 | -93.1741
9.5408 | -93.1747
9.5412 | -93.1753
9.5415 | -93.1759
9.5418 | -93.1766
9.5422 | -93.1771
9.5425 | -93.177
9.5429 | -93.1771
9.5432 | -93.1775
9.5435 | -93.1779
9.5439 | -93.1783
9.5442 | -93.1789
9.5446 | -93.1802
9.5449 | -93.1813
9.5452 | -93.182
9.5456 | -93.1828
9.5459 | -93.1832
9.5463 | -93.1836
9.5466 | -93.1839
9.5469 | -93.1845
9.5473 | -93.185
9.5476 | -93.1853
9.548 | -93.1853
9.5483 | -93.1848
9.5486 | -93.1855
9.549 | -93.1861
9.5493 | -93.1872
9.5497 | -93.1891
9.55 | -93.1905
9.5504 | -93.1913
9.5507 | -93.1918
9.551 | -93.1923
9.5514 | -93.1926
9.5517 | -93.1922
9.5521 | -93.1928
9.5524 | -93.1935
9.5527 | -93.194
9.5531 | -93.195
9.5534 | -93.196
9.5538 | -93.1965
9.5541 | -93.1969
9.5544 | -93.197
9.5548 | -93.1969
9.5551 | -93.1968
9.5555 | -93.1973
9.5558 | -93.1981
9.5561 | -93.1984
9.5565 | -93.1987
9.5568 | -93.1984
9.5572 | -93.1986
9.5575 | -93.1991
9.5578 | -93.1996
9.5582 | -93.1998
9.5585 | -93.2007
9.5589 | -93.2013
9.5592 | -93.2023
9.5595 | -93.2035
9.5599 | -93.2042
9.5602 | -93.2046
9.5606 | -93.2054
9.5609 | -93.2062
9.5612 | -93.207
9.5616 | -93.2089
9.5619 | -93.2103
9.5623 | -93.2104
9.5626 | -93.2112
9.5629 | -93.2115
9.5633 | -93.2123
9.5636 | -93.213
9.564 | -93.2135
9.5643 | -93.2149
9.5647 | -93.2165
9.565 | -93.2172
9.5653 | -93.2179
9.5657 | -93.2185
9.566 | -93.2192
9.5664 | -93.2199
9.5667 | -93.2207
9.567 | -93.2215
9.5674 | -93.222
9.5677 | -93.2224
9.5681 | -93.223
9.5684 | -93.2241
9.5687 | -93.2243
9.5691 | -93.2248
9.5694 | -93.2257
9.5698 | -93.2262
9.5701 | -93.2266
9.5704 | -93.2272
9.5708 | -93.2277
9.5711 | -93.228
9.5715 | -93.229
9.5718 | -93.2299
9.5721 | -93.2301
9.5725 | -93.2306
9.5728 | -93.231
9.5732 | -93.2318
9.5735 | -93.2328
9.5738 | -93.2333
9.5742 | -93.2333
9.5745 | -93.2338
9.5749 | -93.2344
9.5752 | -93.2351
9.5755 | -93.2361
9.5759 | -93.2364
9.5762 | -93.2371
9.5766 | -93.2381
9.5769 | -93.2384
9.5772 | -93.2383
9.5776 | -93.2386
9.5779 | -93.2389
9.5783 | -93.2394
9.5786 | -93.2402
9.579 | -93.241
9.5793 | -93.2415
9.5796 | -93.2423
9.58 | -93.2429
9.5803 | -93.2436
9.5807 | -93.2437
9.581 | -93.2437
9.5813 | -93.2438
9.5817 | -93.245
9.582 | -93.2463
9.5824 | -93.2467
9.5827 | -93.2478
9.583 | -93.2491
9.5834 | -93.2498
9.5837 | -93.2502
9.5841 | -93.251
9.5844 | -93.252
9.5847 | -93.2522
9.5851 | -93.2527
9.5854 | -93.2534
9.5858 | -93.2538
9.5861 | -93.2543
9.5864 | -93.2549
9.5868 | -93.2557
9.5871 | -93.2563
9.5875 | -93.2568
9.5878 | -93.2579
9.5881 | -93.2585
9.5885 | -93.2593
9.5888 | -93.26
9.5892 | -93.2613
9.5895 | -93.2622
9.5898 | -93.2628
9.5902 | -93.2633
9.5905 | -93.2638
9.5909 | -93.2645
9.5912 | -93.265
9.5915 | -93.2653
9.5919 | -93.2659
9.5922 | -93.2664
9.5926 | -93.267
9.5929 | -93.2677
9.5933 | -93.2679
9.5936 | -93.2688
9.5939 | -93.2702
9.5943 | -93.2708
9.5946 | -93.2722
9.595 | -93.2734
9.5953 | -93.274
9.5956 | -93.2746
9.596 | -93.275
9.5963 | -93.2756
9.5967 | -93.2776
9.597 | -93.2793
9.5973 | -93.2804
9.5977 | -93.2815
9.598 | -93.2826
9.5984 | -93.2833
9.5987 | -93.2838
9.599 | -93.2843
9.5994 | -93.2847
9.5997 | -93.2856
9.6001 | -93.2864
9.6004 | -93.2873
9.6007 | -93.2883
9.6011 | -93.2888
9.6014 | -93.2896
9.6018 | -93.2903
9.6021 | -93.2907
9.6024 | -93.291
9.6028 | -93.2918
9.6031 | -93.2927
9.6035 | -93.2934
9.6038 | -93.2941
9.6041 | -93.2948
9.6045 | -93.2955
9.6048 | -93.2964
9.6052 | -93.2969
9.6055 | -93.2974
9.6059 | -93.2978
9.6062 | -93.298
9.6065 | -93.2983
9.6069 | -93.299
9.6072 | -93.3003
9.6076 | -93.3013
9.6079 | -93.3025
9.6082 | -93.3033
9.6086 | -93.3039
9.6089 | -93.3044
9.6093 | -93.305
9.6096 | -93.3058
9.6099 | -93.3064
9.6103 | -93.3067
9.6106 | -93.3072
9.611 | -93.3074
9.6113 | -93.3079
9.6116 | -93.3086
9.612 | -93.3095
9.6123 | -93.3106
9.6127 | -93.3122
9.613 | -93.3137
9.6133 | -93.315
9.6137 | -93.3161
9.614 | -93.3168
9.6144 | -93.3174
9.6147 | -93.3182
9.615 | -93.3189
9.6154 | -93.3197
9.6157 | -93.3204
9.6161 | -93.321
9.6164 | -93.3221
9.6167 | -93.3228
9.6171 | -93.3235
9.6174 | -93.3234
9.6178 | -93.3242
9.6181 | -93.3247
9.6184 | -93.3254
9.6188 | -93.3262
9.6191 | -93.3277
9.6195 | -93.3284
9.6198 | -93.329
9.6202 | -93.3297
9.6205 | -93.3304
9.6208 | -93.3308
9.6212 | -93.3313
9.6215 | -93.3322
9.6219 | -93.3327
9.6222 | -93.3333
9.6225 | -93.3337
9.6229 | -93.3342
9.6232 | -93.3347
9.6236 | -93.3354
9.6239 | -93.3359
9.6242 | -93.3361
9.6246 | -93.3367
9.6249 | -93.3376
9.6253 | -93.3387
9.6256 | -93.34
9.6259 | -93.3412
9.6263 | -93.3426
9.6266 | -93.344
9.627 | -93.3455
9.6273 | -93.3461
9.6276 | -93.3474
9.628 | -93.3483
9.6283 | -93.349
9.6287 | -93.3495
9.629 | -93.3497
9.6293 | -93.3504
9.6297 | -93.3513
9.63 | -93.3522
9.6304 | -93.3525
9.6307 | -93.353
9.631 | -93.3534
9.6314 | -93.3545
9.6317 | -93.3561
9.6321 | -93.3569
9.6324 | -93.357
9.6327 | -93.3589
9.6331 | -93.3599
9.6334 | -93.3605
9.6338 | -93.3613
9.6341 | -93.3628
9.6345 | -93.364
9.6348 | -93.3648
9.6351 | -93.3653
9.6355 | -93.3656
9.6358 | -93.3661
9.6362 | -93.3671
9.6365 | -93.3677
9.6368 | -93.3684
9.6372 | -93.3696
9.6375 | -93.3704
9.6379 | -93.371
9.6382 | -93.3716
9.6385 | -93.3725
9.6389 | -93.3735
9.6392 | -93.3741
9.6396 | -93.3744
9.6399 | -93.3749
9.6402 | -93.3754
9.6406 | -93.3758
9.6409 | -93.3766
9.6413 | -93.3777
9.6416 | -93.3785
9.6419 | -93.3792
9.6423 | -93.38
9.6426 | -93.3801
9.643 | -93.3806
9.6433 | -93.3811
9.6436 | -93.3815
9.644 | -93.3822
9.6443 | -93.3832
9.6447 | -93.384
9.645 | -93.3849
9.6453 | -93.3853
9.6457 | -93.3859
9.646 | -93.3867
9.6464 | -93.3874
9.6467 | -93.3881
9.6471 | -93.3894
9.6474 | -93.3901
9.6477 | -93.391
9.6481 | -93.392
9.6484 | -93.3934
9.6488 | -93.3953
9.6491 | -93.3964
9.6494 | -93.3971
9.6498 | -93.3978
9.6501 | -93.3986
9.6505 | -93.3994
9.6508 | -93.4001
9.6511 | -93.4004
9.6515 | -93.4015
9.6518 | -93.4032
9.6522 | -93.4052
9.6525 | -93.4067
9.6528 | -93.4078
9.6532 | -93.4085
9.6535 | -93.409
9.6539 | -93.4098
9.6542 | -93.4108
9.6545 | -93.4116
9.6549 | -93.4121
9.6552 | -93.4125
9.6556 | -93.4136
9.6559 | -93.415
9.6562 | -93.4159
9.6566 | -93.4171
9.6569 | -93.4184
9.6573 | -93.4197
9.6576 | -93.4206
9.6579 | -93.421
9.6583 | -93.4214
9.6586 | -93.4222
9.659 | -93.423
9.6593 | -93.4239
9.6596 | -93.4248
9.66 | -93.426
9.6603 | -93.427
9.6607 | -93.4278
9.661 | -93.4286
9.6614 | -93.4293
9.6617 | -93.4302
9.662 | -93.4311
9.6624 | -93.4326
9.6627 | -93.4343
9.6631 | -93.4352
9.6634 | -93.4359
9.6637 | -93.437
9.6641 | -93.4382
9.6644 | -93.4391
9.6648 | -93.4399
9.6651 | -93.441
9.6654 | -93.4413
9.6658 | -93.4419
9.6661 | -93.4429
9.6665 | -93.4435
9.6668 | -93.4448
9.6671 | -93.4456
9.6675 | -93.4464
9.6678 | -93.4462
9.6682 | -93.4462
9.6685 | -93.4467
9.6688 | -93.4477
9.6692 | -93.4488
9.6695 | -93.4498
9.6699 | -93.4506
9.6702 | -93.4513
9.6705 | -93.4522
9.6709 | -93.4537
9.6712 | -93.4557
9.6716 | -93.4571
9.6719 | -93.4579
9.6722 | -93.4588
9.6726 | -93.4597
9.6729 | -93.4605
9.6733 | -93.4606
9.6736 | -93.4604
9.6739 | -93.4611
9.6743 | -93.4621
9.6746 | -93.4628
9.675 | -93.4634
9.6753 | -93.4644
9.6757 | -93.465
9.676 | -93.4655
9.6763 | -93.4665
9.6767 | -93.4666
9.677 | -93.4671
9.6774 | -93.4678
9.6777 | -93.4674
9.678 | -93.4672
9.6784 | -93.4677
9.6787 | -93.4686
9.6791 | -93.4694
9.6794 | -93.4703
9.6797 | -93.4713
9.6801 | -93.4728
9.6804 | -93.4739
9.6808 | -93.4746
9.6811 | -93.4752
9.6814 | -93.4762
9.6818 | -93.4772
9.6821 | -93.4779
9.6825 | -93.4787
9.6828 | -93.4797
9.6831 | -93.4806
9.6835 | -93.4814
9.6838 | -93.4823
9.6842 | -93.4833
9.6845 | -93.4842
9.6848 | -93.4851
9.6852 | -93.4862
9.6855 | -93.487
9.6859 | -93.4876
9.6862 | -93.4887
9.6865 | -93.4896
9.6869 | -93.4903
9.6872 | -93.4911
9.6876 | -93.4914
9.6879 | -93.4922
9.6882 | -93.4933
9.6886 | -93.494
9.6889 | -93.4947
9.6893 | -93.4956
9.6896 | -93.4964
9.69 | -93.4972
9.6903 | -93.4983
9.6906 | -93.4987
9.691 | -93.4999
9.6913 | -93.5012
9.6917 | -93.5022
9.692 | -93.5031
9.6923 | -93.5036
9.6927 | -93.504
9.693 | -93.5049
9.6934 | -93.5061
9.6937 | -93.5074
9.694 | -93.5085
9.6944 | -93.5096
9.6947 | -93.5108
9.6951 | -93.5121
9.6954 | -93.513
9.6957 | -93.5136
9.6961 | -93.5144
9.6964 | -93.5154
9.6968 | -93.5167
9.6971 | -93.5173
9.6974 | -93.5176
9.6978 | -93.5184
9.6981 | -93.5198
9.6985 | -93.5209
9.6988 | -93.5216
9.6991 | -93.522
9.6995 | -93.5229
9.6998 | -93.5237
9.7002 | -93.5247
9.7005 | -93.5257
9.7008 | -93.5265
9.7012 | -93.5277
9.7015 | -93.5291
9.7019 | -93.5302
9.7022 | -93.5311
9.7026 | -93.5323
9.7029 | -93.5332
9.7032 | -93.534
9.7036 | -93.535
9.7039 | -93.5357
9.7043 | -93.5369
9.7046 | -93.5383
9.7049 | -93.5395
9.7053 | -93.5405
9.7056 | -93.5413
9.706 | -93.5426
9.7063 | -93.5437
9.7066 | -93.5447
9.707 | -93.5455
9.7073 | -93.5464
9.7077 | -93.5475
9.708 | -93.5484
9.7083 | -93.5496
9.7087 | -93.5503
9.709 | -93.5513
9.7094 | -93.5522
9.7097 | -93.5535
9.71 | -93.554
9.7104 | -93.5549
9.7107 | -93.5558
9.7111 | -93.5562
9.7114 | -93.5571
9.7117 | -93.558
9.7121 | -93.5591
9.7124 | -93.5606
9.7128 | -93.5622
9.7131 | -93.5638
9.7134 | -93.5647
9.7138 | -93.5658
9.7141 | -93.5668
9.7145 | -93.5674
9.7148 | -93.5687
9.7151 | -93.5707
9.7155 | -93.5718
9.7158 | -93.5724
9.7162 | -93.573
9.7165 | -93.5745
9.7169 | -93.5755
9.7172 | -93.5769
9.7175 | -93.5795
9.7179 | -93.5818
9.7182 | -93.583
9.7186 | -93.584
9.7189 | -93.585
9.7192 | -93.5858
9.7196 | -93.5866
9.7199 | -93.5875
9.7203 | -93.5882
9.7206 | -93.5879
9.7209 | -93.5876
9.7213 | -93.5884
9.7216 | -93.5892
9.722 | -93.5901
9.7223 | -93.5913
9.7226 | -93.5921
9.723 | -93.593
9.7233 | -93.5947
9.7237 | -93.5957
9.724 | -93.597
9.7243 | -93.5977
9.7247 | -93.5984
9.725 | -93.5994
9.7254 | -93.6007
9.7257 | -93.6017
9.726 | -93.6023
9.7264 | -93.6028
9.7267 | -93.6033
9.7271 | -93.6034
9.7274 | -93.604
9.7277 | -93.605
9.7281 | -93.6062
9.7284 | -93.6078
9.7288 | -93.6087
9.7291 | -93.6091
9.7294 | -93.61
9.7298 | -93.6109
9.7301 | -93.6123
9.7305 | -93.6139
9.7308 | -93.6144
9.7312 | -93.6149
9.7315 | -93.6162
9.7318 | -93.6173
9.7322 | -93.6181
9.7325 | -93.6184
9.7329 | -93.6187
9.7332 | -93.6197
9.7335 | -93.6208
9.7339 | -93.6216
9.7342 | -93.6222
9.7346 | -93.6234
9.7349 | -93.6243
9.7352 | -93.6253
9.7356 | -93.6262
9.7359 | -93.6271
9.7363 | -93.6279
9.7366 | -93.6289
9.7369 | -93.6301
9.7373 | -93.6313
9.7376 | -93.632
9.738 | -93.6329
9.7383 | -93.6338
9.7386 | -93.6343
9.739 | -93.6354
9.7393 | -93.6362
9.7397 | -93.6369
9.74 | -93.6378
9.7403 | -93.6385
9.7407 | -93.6394
9.741 | -93.6402
9.7414 | -93.6411
9.7417 | -93.6427
9.742 | -93.6436
9.7424 | -93.6447
9.7427 | -93.6458
9.7431 | -93.6464
9.7434 | -93.6473
9.7437 | -93.6486
9.7441 | -93.6494
9.7444 | -93.6504
9.7448 | -93.6516
9.7451 | -93.6521
9.7455 | -93.6529
9.7458 | -93.6534
9.7461 | -93.6542
9.7465 | -93.6559
9.7468 | -93.6577
9.7472 | -93.6595
9.7475 | -93.6605
9.7478 | -93.6612
9.7482 | -93.6625
9.7485 | -93.6638
9.7489 | -93.6651
9.7492 | -93.6661
9.7495 | -93.6668
9.7499 | -93.6675
9.7502 | -93.6682
9.7506 | -93.6691
9.7509 | -93.6699
9.7512 | -93.671
9.7516 | -93.6723
9.7519 | -93.6735
9.7523 | -93.6742
9.7526 | -93.675
9.7529 | -93.6759
9.7533 | -93.6765
9.7536 | -93.677
9.754 | -93.678
9.7543 | -93.6789
9.7546 | -93.6794
9.755 | -93.6793
9.7553 | -93.6794
9.7557 | -93.6809
9.756 | -93.6825
9.7563 | -93.6831
9.7567 | -93.684
9.757 | -93.6845
9.7574 | -93.6851
9.7577 | -93.6862
9.758 | -93.6869
9.7584 | -93.6876
9.7587 | -93.6885
9.7591 | -93.6893
9.7594 | -93.6899
9.7598 | -93.6907
9.7601 | -93.6918
9.7604 | -93.6931
9.7608 | -93.6944
9.7611 | -93.6955
9.7615 | -93.6964
9.7618 | -93.6975
9.7621 | -93.6983
9.7625 | -93.6994
9.7628 | -93.7001
9.7632 | -93.7009
9.7635 | -93.7016
9.7638 | -93.7017
9.7642 | -93.7023
9.7645 | -93.7033
9.7649 | -93.7042
9.7652 | -93.7049
9.7655 | -93.7054
9.7659 | -93.7064
9.7662 | -93.7078
9.7666 | -93.7087
9.7669 | -93.7097
9.7672 | -93.7108
9.7676 | -93.7119
9.7679 | -93.7131
9.7683 | -93.7141
9.7686 | -93.7148
9.7689 | -93.7159
9.7693 | -93.7173
9.7696 | -93.7184
9.77 | -93.7194
9.7703 | -93.7202
9.7706 | -93.7209
9.771 | -93.7221
9.7713 | -93.723
9.7717 | -93.724
9.772 | -93.7249
9.7723 | -93.7261
9.7727 | -93.7269
9.773 | -93.7279
9.7734 | -93.729
9.7737 | -93.7303
9.7741 | -93.7313
9.7744 | -93.7316
9.7747 | -93.7322
9.7751 | -93.733
9.7754 | -93.734
9.7758 | -93.7349
9.7761 | -93.7358
9.7764 | -93.7369
9.7768 | -93.7381
9.7771 | -93.739
9.7775 | -93.74
9.7778 | -93.7409
9.7781 | -93.7415
9.7785 | -93.7421
9.7788 | -93.7426
9.7792 | -93.7435
9.7795 | -93.745
9.7798 | -93.7473
9.7802 | -93.7484
9.7805 | -93.7488
9.7809 | -93.7509
9.7812 | -93.7527
9.7815 | -93.7539
9.7819 | -93.7548
9.7822 | -93.7556
9.7826 | -93.7565
9.7829 | -93.757
9.7832 | -93.7581
9.7836 | -93.759
9.7839 | -93.7598
9.7843 | -93.7605
9.7846 | -93.7598
9.7849 | -93.7598
9.7853 | -93.7605
9.7856 | -93.7609
9.786 | -93.7623
9.7863 | -93.7626
9.7866 | -93.7628
9.787 | -93.7641
9.7873 | -93.7657
9.7877 | -93.7666
9.788 | -93.7675
9.7884 | -93.7683
9.7887 | -93.7689
9.789 | -93.7693
9.7894 | -93.7701
9.7897 | -93.7706
9.7901 | -93.771
9.7904 | -93.7721
9.7907 | -93.7733
9.7911 | -93.7743
9.7914 | -93.7759
9.7918 | -93.777
9.7921 | -93.778
9.7924 | -93.7794
9.7928 | -93.7796
9.7931 | -93.7801
9.7935 | -93.7811
9.7938 | -93.7815
9.7941 | -93.7821
9.7945 | -93.7829
9.7948 | -93.7838
9.7952 | -93.7846
9.7955 | -93.7855
9.7958 | -93.7861
9.7962 | -93.7864
9.7965 | -93.7867
9.7969 | -93.7876
9.7972 | -93.7883
9.7975 | -93.7893
9.7979 | -93.7904
9.7982 | -93.791
9.7986 | -93.7918
9.7989 | -93.7928
9.7992 | -93.7937
9.7996 | -93.795
9.7999 | -93.7961
9.8003 | -93.7983
9.8006 | -93.8009
9.801 | -93.8029
9.8013 | -93.8037
9.8016 | -93.8046
9.802 | -93.8051
9.8023 | -93.8055
9.8027 | -93.8071
9.803 | -93.8093
9.8033 | -93.8104
9.8037 | -93.8111
9.804 | -93.8117
9.8044 | -93.8134
9.8047 | -93.8148
9.805 | -93.8156
9.8054 | -93.8164
9.8057 | -93.817
9.8061 | -93.8177
9.8064 | -93.8184
9.8067 | -93.8192
9.8071 | -93.8199
9.8074 | -93.8209
9.8078 | -93.8221
9.8081 | -93.8229
9.8084 | -93.8235
9.8088 | -93.8241
9.8091 | -93.8249
9.8095 | -93.826
9.8098 | -93.8275
9.8101 | -93.8287
9.8105 | -93.8296
9.8108 | -93.8307
9.8112 | -93.8319
9.8115 | -93.8322
9.8118 | -93.8326
9.8122 | -93.8332
9.8125 | -93.8339
9.8129 | -93.8352
9.8132 | -93.8364
9.8135 | -93.8371
9.8139 | -93.8381
9.8142 | -93.8391
9.8146 | -93.8403
9.8149 | -93.841
9.8153 | -93.8417
9.8156 | -93.8425
9.8159 | -93.8432
9.8163 | -93.844
9.8166 | -93.845
9.817 | -93.8465
9.8173 | -93.8482
9.8176 | -93.8492
9.818 | -93.8499
9.8183 | -93.8506
9.8187 | -93.8515
9.819 | -93.8525
9.8193 | -93.8534
9.8197 | -93.8542
9.82 | -93.8551
9.8204 | -93.8558
9.8207 | -93.8566
9.821 | -93.8574
9.8214 | -93.8582
9.8217 | -93.8589
9.8221 | -93.8597
9.8224 | -93.8604
9.8227 | -93.8621
9.8231 | -93.8635
9.8234 | -93.8644
9.8238 | -93.8645
9.8241 | -93.8661
9.8244 | -93.8671
9.8248 | -93.8676
9.8251 | -93.8678
9.8255 | -93.8685
9.8258 | -93.8693
9.8261 | -93.87
9.8265 | -93.8717
9.8268 | -93.8734
9.8272 | -93.8745
9.8275 | -93.8753
9.8278 | -93.876
9.8282 | -93.8774
9.8285 | -93.8781
9.8289 | -93.8787
9.8292 | -93.8794
9.8296 | -93.8807
9.8299 | -93.8817
9.8302 | -93.8825
9.8306 | -93.8833
9.8309 | -93.8845
9.8313 | -93.8853
9.8316 | -93.8858
9.8319 | -93.8864
9.8323 | -93.8873
9.8326 | -93.8882
9.833 | -93.8885
9.8333 | -93.8893
9.8336 | -93.89
9.834 | -93.8909
9.8343 | -93.8917
9.8347 | -93.8923
9.835 | -93.8929
9.8353 | -93.8936
9.8357 | -93.8943
9.836 | -93.895
9.8364 | -93.8958
9.8367 | -93.8967
9.837 | -93.8974
9.8374 | -93.8979
9.8377 | -93.8986
9.8381 | -93.8999
9.8384 | -93.901
9.8387 | -93.9006
9.8391 | -93.9007
9.8394 | -93.9017
9.8398 | -93.9022
9.8401 | -93.9028
9.8404 | -93.9032
9.8408 | -93.9037
9.8411 | -93.9042
9.8415 | -93.9049
9.8418 | -93.9057
9.8422 | -93.9064
9.8425 | -93.9072
9.8428 | -93.9082
9.8432 | -93.909
9.8435 | -93.91
9.8439 | -93.9112
9.8442 | -93.9123
9.8445 | -93.9128
9.8449 | -93.9134
9.8452 | -93.9142
9.8456 | -93.9138
9.8459 | -93.9134
9.8462 | -93.9135
9.8466 | -93.9141
9.8469 | -93.9148
9.8473 | -93.9155
9.8476 | -93.9161
9.8479 | -93.9167
9.8483 | -93.9174
9.8486 | -93.9183
9.849 | -93.919
9.8493 | -93.9195
9.8496 | -93.9201
9.85 | -93.9203
9.8503 | -93.9204
9.8507 | -93.921
9.851 | -93.922
9.8513 | -93.9232
9.8517 | -93.9239
9.852 | -93.9246
9.8524 | -93.9256
9.8527 | -93.9262
9.853 | -93.9268
9.8534 | -93.9278
9.8537 | -93.9284
9.8541 | -93.9289
9.8544 | -93.9297
9.8547 | -93.9303
9.8551 | -93.9302
9.8554 | -93.9302
9.8558 | -93.9311
9.8561 | -93.9319
9.8565 | -93.9326
9.8568 | -93.9334
9.8571 | -93.9344
9.8575 | -93.9349
9.8578 | -93.9358
9.8582 | -93.9369
9.8585 | -93.9376
9.8588 | -93.9383
9.8592 | -93.9394
9.8595 | -93.9403
9.8599 | -93.9409
9.8602 | -93.9416
9.8605 | -93.9427
9.8609 | -93.9434
9.8612 | -93.944
9.8616 | -93.9444
9.8619 | -93.945
9.8622 | -93.9459
9.8626 | -93.9464
9.8629 | -93.9464
9.8633 | -93.9469
9.8636 | -93.9476
9.8639 | -93.9481
9.8643 | -93.9488
9.8646 | -93.9493
9.865 | -93.9502
9.8653 | -93.9512
9.8656 | -93.9526
9.866 | -93.9529
9.8663 | -93.9536
9.8667 | -93.9544
9.867 | -93.9551
9.8673 | -93.9555
9.8677 | -93.956
9.868 | -93.9569
9.8684 | -93.9575
9.8687 | -93.9581
9.869 | -93.9588
9.8694 | -93.9598
9.8697 | -93.9609
9.8701 | -93.9617
9.8704 | -93.9627
9.8708 | -93.9638
9.8711 | -93.9656
9.8714 | -93.9669
9.8718 | -93.9678
9.8721 | -93.9683
9.8725 | -93.9684
9.8728 | -93.9692
9.8731 | -93.9698
9.8735 | -93.9707
9.8738 | -93.9714
9.8742 | -93.9722
9.8745 | -93.9736
9.8748 | -93.9745
9.8752 | -93.9753
9.8755 | -93.9758
9.8759 | -93.9762
9.8762 | -93.9768
9.8765 | -93.9774
9.8769 | -93.9781
9.8772 | -93.9787
9.8776 | -93.9791
9.8779 | -93.9797
9.8782 | -93.9802
9.8786 | -93.9809
9.8789 | -93.9813
9.8793 | -93.9824
9.8796 | -93.983
9.8799 | -93.9836
9.8803 | -93.9841
9.8806 | -93.9841
9.881 | -93.9846
9.8813 | -93.986
9.8816 | -93.9873
9.882 | -93.9887
9.8823 | -93.9896
9.8827 | -93.9905
9.883 | -93.9909
9.8833 | -93.9915
9.8837 | -93.9924
9.884 | -93.9937
9.8844 | -93.9948
9.8847 | -93.9953
9.8851 | -93.9961
9.8854 | -93.997
9.8857 | -93.9984
9.8861 | -93.9994
9.8864 | -93.9999
9.8868 | -94.0001
9.8871 | -94.001
9.8874 | -94.0017
9.8878 | -94.0022
9.8881 | -94.0028
9.8885 | -94.0031
9.8888 | -94.0037
9.8891 | -94.0041
9.8895 | -94.0044
9.8898 | -94.005
9.8902 | -94.0057
9.8905 | -94.0068
9.8908 | -94.0075
9.8912 | -94.0082
9.8915 | -94.0087
9.8919 | -94.0095
9.8922 | -94.0099
9.8925 | -94.0107
9.8929 | -94.0111
9.8932 | -94.0116
9.8936 | -94.0126
9.8939 | -94.0133
9.8942 | -94.014
9.8946 | -94.0145
9.8949 | -94.0152
9.8953 | -94.0162
9.8956 | -94.0173
9.8959 | -94.0179
9.8963 | -94.0184
9.8966 | -94.0194
9.897 | -94.0202
9.8973 | -94.0208
9.8976 | -94.0218
9.898 | -94.0228
9.8983 | -94.0233
9.8987 | -94.0239
9.899 | -94.0245
9.8994 | -94.0253
9.8997 | -94.026
9.9 | -94.0264
9.9004 | -94.0271
9.9007 | -94.0278
9.9011 | -94.0285
9.9014 | -94.029
9.9017 | -94.03
9.9021 | -94.0311
9.9024 | -94.0318
9.9028 | -94.0325
9.9031 | -94.0331
9.9034 | -94.0338
9.9038 | -94.0344
9.9041 | -94.0349
9.9045 | -94.0359
9.9048 | -94.0368
9.9051 | -94.0382
9.9055 | -94.0397
9.9058 | -94.0409
9.9062 | -94.0418
9.9065 | -94.0421
9.9068 | -94.0425
9.9072 | -94.043
9.9075 | -94.0436
9.9079 | -94.0445
9.9082 | -94.0453
9.9085 | -94.0466
9.9089 | -94.0476
9.9092 | -94.0488
9.9096 | -94.0495
9.9099 | -94.0501
9.9102 | -94.051
9.9106 | -94.0514
9.9109 | -94.0517
9.9113 | -94.0525
9.9116 | -94.0531
9.9119 | -94.0537
9.9123 | -94.054
9.9126 | -94.0546
9.913 | -94.0551
9.9133 | -94.0556
9.9137 | -94.0561
9.914 | -94.0567
9.9143 | -94.0572
9.9147 | -94.0578
9.915 | -94.0583
9.9154 | -94.0587
9.9157 | -94.0592
9.916 | -94.0593
9.9164 | -94.0599
9.9167 | -94.0605
9.9171 | -94.0612
9.9174 | -94.0615
9.9177 | -94.0615
9.9181 | -94.0623
9.9184 | -94.0632
9.9188 | -94.0638
9.9191 | -94.0642
9.9194 | -94.0655
9.9198 | -94.066
9.9201 | -94.0667
9.9205 | -94.0675
9.9208 | -94.0681
9.9211 | -94.0686
9.9215 | -94.0695
9.9218 | -94.0699
9.9222 | -94.0703
9.9225 | -94.0704
9.9228 | -94.0711
9.9232 | -94.0717
9.9235 | -94.0729
9.9239 | -94.0751
9.9242 | -94.0766
9.9245 | -94.0773
9.9249 | -94.0776
9.9252 | -94.0781
9.9256 | -94.0786
9.9259 | -94.0791
9.9262 | -94.0795
9.9266 | -94.08
9.9269 | -94.0805
9.9273 | -94.0817
9.9276 | -94.0834
9.928 | -94.0849
9.9283 | -94.0855
9.9286 | -94.0861
9.929 | -94.0868
9.9293 | -94.0876
9.9297 | -94.0883
9.93 | -94.0888
9.9303 | -94.0893
9.9307 | -94.0892
9.931 | -94.0886
9.9314 | -94.0891
9.9317 | -94.0896
9.932 | -94.0903
9.9324 | -94.0909
9.9327 | -94.0917
9.9331 | -94.0919
9.9334 | -94.0924
9.9337 | -94.0929
9.9341 | -94.0933
9.9344 | -94.0937
9.9348 | -94.0941
9.9351 | -94.0946
9.9354 | -94.0951
9.9358 | -94.0954
9.9361 | -94.096
9.9365 | -94.0967
9.9368 | -94.0977
9.9371 | -94.0983
9.9375 | -94.0985
9.9378 | -94.0992
9.9382 | -94.0999
9.9385 | -94.1007
9.9388 | -94.1018
9.9392 | -94.1023
9.9395 | -94.1028
9.9399 | -94.1033
9.9402 | -94.104
9.9406 | -94.105
9.9409 | -94.106
9.9412 | -94.1068
9.9416 | -94.1059
9.9419 | -94.1064
9.9423 | -94.1074
9.9426 | -94.1085
9.9429 | -94.109
9.9433 | -94.1097
9.9436 | -94.1102
9.944 | -94.1106
9.9443 | -94.1106
9.9446 | -94.1114
9.945 | -94.1117
9.9453 | -94.1121
9.9457 | -94.1136
9.946 | -94.1149
9.9463 | -94.1157
9.9467 | -94.1166
9.947 | -94.1181
9.9474 | -94.1199
9.9477 | -94.1207
9.948 | -94.1213
9.9484 | -94.1261
9.9487 | -94.1301
9.9491 | -94.1307
9.9494 | -94.131
9.9497 | -94.1318
9.9501 | -94.1323
9.9504 | -94.1327
9.9508 | -94.1333
9.9511 | -94.1334
9.9514 | -94.1341
9.9518 | -94.1341
9.9521 | -94.1345
9.9525 | -94.135
9.9528 | -94.1358
9.9531 | -94.1365
9.9535 | -94.1372
9.9538 | -94.1387
9.9542 | -94.1397
9.9545 | -94.14
9.9549 | -94.1409
9.9552 | -94.1419
9.9555 | -94.1422
9.9559 | -94.1429
9.9562 | -94.1436
9.9566 | -94.144
9.9569 | -94.1446
9.9572 | -94.1451
9.9576 | -94.1452
9.9579 | -94.1456
9.9583 | -94.146
9.9586 | -94.1462
9.9589 | -94.1468
9.9593 | -94.1476
9.9596 | -94.1482
9.96 | -94.1487
9.9603 | -94.1487
9.9606 | -94.149
9.961 | -94.1493
9.9613 | -94.1502
9.9617 | -94.1508
9.962 | -94.1511
9.9623 | -94.1513
9.9627 | -94.1519
9.963 | -94.1527
9.9634 | -94.154
9.9637 | -94.1547
9.964 | -94.1552
9.9644 | -94.1555
9.9647 | -94.1556
9.9651 | -94.1562
9.9654 | -94.1568
9.9657 | -94.1574
9.9661 | -94.1582
9.9664 | -94.1585
9.9668 | -94.1591
9.9671 | -94.1604
9.9674 | -94.1614
9.9678 | -94.1617
9.9681 | -94.1619
9.9685 | -94.1622
9.9688 | -94.1626
9.9692 | -94.1631
9.9695 | -94.1637
9.9698 | -94.1642
9.9702 | -94.1647
9.9705 | -94.1653
9.9709 | -94.1662
9.9712 | -94.167
9.9715 | -94.1681
9.9719 | -94.1687
9.9722 | -94.1687
9.9726 | -94.1684
9.9729 | -94.1687
9.9732 | -94.1692
9.9736 | -94.1696
9.9739 | -94.1701
9.9743 | -94.1707
9.9746 | -94.1712
9.9749 | -94.172
9.9753 | -94.1723
9.9756 | -94.1727
9.976 | -94.1734
9.9763 | -94.1741
9.9766 | -94.1747
9.977 | -94.1753
9.9773 | -94.1756
9.9777 | -94.176
9.978 | -94.1765
9.9783 | -94.1768
9.9787 | -94.1773
9.979 | -94.1778
9.9794 | -94.1786
9.9797 | -94.1795
9.98 | -94.1798
9.9804 | -94.1801
9.9807 | -94.1807
9.9811 | -94.1815
9.9814 | -94.1821
9.9817 | -94.1825
9.9821 | -94.1826
9.9824 | -94.1833
9.9828 | -94.1845
9.9831 | -94.1854
9.9835 | -94.1865
9.9838 | -94.1877
9.9841 | -94.1885
9.9845 | -94.1894
9.9848 | -94.1908
9.9852 | -94.1919
9.9855 | -94.1925
9.9858 | -94.193
9.9862 | -94.1935
9.9865 | -94.1941
9.9869 | -94.1947
9.9872 | -94.1957
9.9875 | -94.197
9.9879 | -94.1979
9.9882 | -94.1983
9.9886 | -94.1994
9.9889 | -94.2001
9.9892 | -94.2004
9.9896 | -94.2006
9.9899 | -94.2011
9.9903 | -94.2016
9.9906 | -94.2019
9.9909 | -94.2022
9.9913 | -94.2025
9.9916 | -94.2028
9.992 | -94.2031
9.9923 | -94.2038
9.9926 | -94.2045
9.993 | -94.2055
9.9933 | -94.2063
9.9937 | -94.2071
9.994 | -94.2079
9.9943 | -94.2087
9.9947 | -94.2093
9.995 | -94.2097
9.9954 | -94.2099
9.9957 | -94.2108
9.9961 | -94.2119
9.9964 | -94.2133
9.9967 | -94.2139
9.9971 | -94.2145
9.9974 | -94.2149
9.9978 | -94.2155
9.9981 | -94.2162
9.9984 | -94.2169
9.9988 | -94.2175
9.9991 | -94.2182
9.9995 | -94.2193
9.9998 | -94.2199
10.0001 | -94.2208
10.0005 | -94.2218
10.0008 | -94.2219
10.0012 | -94.2221
10.0015 | -94.2229
10.0018 | -94.223
10.0022 | -94.2231
10.0025 | -94.2237
10.0029 | -94.2241
10.0032 | -94.2238
10.0035 | -94.2235
10.0039 | -94.2234
10.0042 | -94.2236
10.0046 | -94.2241
10.0049 | -94.225
10.0052 | -94.2259
10.0056 | -94.2265
10.0059 | -94.227
10.0063 | -94.2273
10.0066 | -94.2278
10.0069 | -94.2285
10.0073 | -94.2295
10.0076 | -94.2312
10.008 | -94.2328
10.0083 | -94.2338
10.0086 | -94.2349
10.009 | -94.2368
10.0093 | -94.2381
10.0097 | -94.2389
10.01 | -94.2398
10.0104 | -94.2406
10.0107 | -94.2413
10.011 | -94.2419
10.0114 | -94.2425
10.0117 | -94.243
10.0121 | -94.2438
10.0124 | -94.2452
10.0127 | -94.2465
10.0131 | -94.2472
10.0134 | -94.2479
10.0138 | -94.2484
10.0141 | -94.2487
10.0144 | -94.2489
10.0148 | -94.2493
10.0151 | -94.2501
10.0155 | -94.2513
10.0158 | -94.253
10.0161 | -94.254
10.0165 | -94.2547
10.0168 | -94.255
10.0172 | -94.2552
10.0175 | -94.2565
10.0178 | -94.2579
10.0182 | -94.2584
10.0185 | -94.2589
10.0189 | -94.26
10.0192 | -94.2608
10.0195 | -94.2611
10.0199 | -94.2613
10.0202 | -94.2623
10.0206 | -94.2624
10.0209 | -94.2633
10.0212 | -94.2655
10.0216 | -94.266
10.0219 | -94.2668
10.0223 | -94.2673
10.0226 | -94.2673
10.0229 | -94.2676
10.0233 | -94.2681
10.0236 | -94.2683
10.024 | -94.2687
10.0243 | -94.2693
10.0247 | -94.2699
10.025 | -94.271
10.0253 | -94.2724
10.0257 | -94.2733
10.026 | -94.2741
10.0264 | -94.2747
10.0267 | -94.2754
10.027 | -94.2764
10.0274 | -94.2769
10.0277 | -94.2771
10.0281 | -94.2768
10.0284 | -94.2765
10.0287 | -94.2769
10.0291 | -94.2771
10.0294 | -94.2774
10.0298 | -94.2779
10.0301 | -94.2787
10.0304 | -94.2788
10.0308 | -94.2793
10.0311 | -94.2797
10.0315 | -94.2802
10.0318 | -94.2809
10.0321 | -94.2818
10.0325 | -94.2825
10.0328 | -94.2831
10.0332 | -94.2839
10.0335 | -94.2851
10.0338 | -94.286
10.0342 | -94.2868
10.0345 | -94.2881
10.0349 | -94.2892
10.0352 | -94.2902
10.0355 | -94.2913
10.0359 | -94.292
10.0362 | -94.2922
10.0366 | -94.2927
10.0369 | -94.2932
10.0372 | -94.2939
10.0376 | -94.2947
10.0379 | -94.296
10.0383 | -94.2971
10.0386 | -94.298
10.039 | -94.2987
10.0393 | -94.2991
10.0396 | -94.2998
10.04 | -94.3005
10.0403 | -94.3005
10.0407 | -94.3008
10.041 | -94.3015
10.0413 | -94.3023
10.0417 | -94.3041
10.042 | -94.3046
10.0424 | -94.3041
10.0427 | -94.3038
10.043 | -94.3046
10.0434 | -94.3051
10.0437 | -94.3054
10.0441 | -94.3067
10.0444 | -94.3077
10.0447 | -94.3081
10.0451 | -94.3086
10.0454 | -94.3092
10.0458 | -94.3103
10.0461 | -94.3111
10.0464 | -94.3115
10.0468 | -94.3116
10.0471 | -94.3118
10.0475 | -94.3122
10.0478 | -94.3125
10.0481 | -94.3131
10.0485 | -94.3133
10.0488 | -94.314
10.0492 | -94.3144
10.0495 | -94.315
10.0498 | -94.3161
10.0502 | -94.3167
10.0505 | -94.3172
10.0509 | -94.3175
10.0512 | -94.3178
10.0516 | -94.3184
10.0519 | -94.3192
10.0522 | -94.3195
10.0526 | -94.3199
10.0529 | -94.3206
10.0533 | -94.3215
10.0536 | -94.3224
10.0539 | -94.3235
10.0543 | -94.3245
10.0546 | -94.3251
10.055 | -94.3259
10.0553 | -94.326
10.0556 | -94.3263
10.056 | -94.3268
10.0563 | -94.3274
10.0567 | -94.3277
10.057 | -94.3283
10.0573 | -94.3289
10.0577 | -94.3292
10.058 | -94.3294
10.0584 | -94.3294
10.0587 | -94.3296
10.059 | -94.33
10.0594 | -94.3311
10.0597 | -94.3317
10.0601 | -94.3322
10.0604 | -94.3326
10.0607 | -94.3335
10.0611 | -94.3347
10.0614 | -94.3357
10.0618 | -94.3363
10.0621 | -94.337
10.0624 | -94.3381
10.0628 | -94.3395
10.0631 | -94.3406
10.0635 | -94.3415
10.0638 | -94.3423
10.0641 | -94.3429
10.0645 | -94.3435
10.0648 | -94.3441
10.0652 | -94.3449
10.0655 | -94.3456
10.0658 | -94.3464
10.0662 | -94.3473
10.0665 | -94.3487
10.0669 | -94.3498
10.0672 | -94.3504
10.0676 | -94.3514
10.0679 | -94.3523
10.0682 | -94.3531
10.0686 | -94.3551
10.0689 | -94.3563
10.0693 | -94.3573
10.0696 | -94.3579
10.0699 | -94.3585
10.0703 | -94.3588
10.0706 | -94.3594
10.071 | -94.36
10.0713 | -94.3606
10.0716 | -94.3603
10.072 | -94.3602
10.0723 | -94.3607
10.0727 | -94.3612
10.073 | -94.3617
10.0733 | -94.3622
10.0737 | -94.3628
10.074 | -94.3638
10.0744 | -94.3648
10.0747 | -94.3654
10.075 | -94.3661
10.0754 | -94.3668
10.0757 | -94.3673
10.0761 | -94.3678
10.0764 | -94.3684
10.0767 | -94.3695
10.0771 | -94.3707
10.0774 | -94.3723
10.0778 | -94.3732
10.0781 | -94.3741
10.0784 | -94.374
10.0788 | -94.3743
10.0791 | -94.3755
10.0795 | -94.3762
10.0798 | -94.3769
10.0802 | -94.3775
10.0805 | -94.3781
10.0808 | -94.3779
10.0812 | -94.3781
10.0815 | -94.3788
10.0819 | -94.3798
10.0822 | -94.3807
10.0825 | -94.3812
10.0829 | -94.3817
10.0832 | -94.3822
10.0836 | -94.3828
10.0839 | -94.3834
10.0842 | -94.3842
10.0846 | -94.3854
10.0849 | -94.3864
10.0853 | -94.3866
10.0856 | -94.3869
10.0859 | -94.3875
10.0863 | -94.388
10.0866 | -94.3882
10.087 | -94.3887
10.0873 | -94.3895
10.0876 | -94.3902
10.088 | -94.3908
10.0883 | -94.3918
10.0887 | -94.3925
10.089 | -94.3936
10.0893 | -94.3944
10.0897 | -94.3949
10.09 | -94.3954
10.0904 | -94.396
10.0907 | -94.3962
10.091 | -94.3965
10.0914 | -94.3978
10.0917 | -94.3985
10.0921 | -94.3991
10.0924 | -94.3998
10.0927 | -94.4005
10.0931 | -94.4016
10.0934 | -94.4024
10.0938 | -94.4039
10.0941 | -94.4054
10.0945 | -94.406
10.0948 | -94.4065
10.0951 | -94.4073
10.0955 | -94.4085
10.0958 | -94.4093
10.0962 | -94.4097
10.0965 | -94.41
10.0968 | -94.4104
10.0972 | -94.4108
10.0975 | -94.4114
10.0979 | -94.412
10.0982 | -94.4128
10.0985 | -94.4133
10.0989 | -94.4145
10.0992 | -94.4151
10.0996 | -94.4158
10.0999 | -94.4165
10.1002 | -94.4171
10.1006 | -94.4179
10.1009 | -94.4182
10.1013 | -94.4188
10.1016 | -94.4197
10.1019 | -94.4207
10.1023 | -94.4214
10.1026 | -94.422
10.103 | -94.4226
10.1033 | -94.4236
10.1036 | -94.4246
10.104 | -94.4253
10.1043 | -94.4261
10.1047 | -94.4274
10.105 | -94.4277
10.1053 | -94.4284
10.1057 | -94.4289
10.106 | -94.4298
10.1064 | -94.431
10.1067 | -94.4319
10.107 | -94.4327
10.1074 | -94.4338
10.1077 | -94.4347
10.1081 | -94.4357
10.1084 | -94.4366
10.1088 | -94.4376
10.1091 | -94.438
10.1094 | -94.4386
10.1098 | -94.4396
10.1101 | -94.4404
10.1105 | -94.4416
10.1108 | -94.4427
10.1111 | -94.4435
10.1115 | -94.4444
10.1118 | -94.4451
10.1122 | -94.4461
10.1125 | -94.4472
10.1128 | -94.4484
10.1132 | -94.4482
10.1135 | -94.4482
10.1139 | -94.4495
10.1142 | -94.45
10.1145 | -94.4511
10.1149 | -94.4519
10.1152 | -94.4529
10.1156 | -94.4539
10.1159 | -94.4543
10.1162 | -94.4551
10.1166 | -94.4563
10.1169 | -94.4574
10.1173 | -94.4582
10.1176 | -94.4587
10.1179 | -94.4594
10.1183 | -94.4605
10.1186 | -94.4617
10.119 | -94.4624
10.1193 | -94.4631
10.1196 | -94.4642
10.12 | -94.4656
10.1203 | -94.4669
10.1207 | -94.468
10.121 | -94.4686
10.1213 | -94.4692
10.1217 | -94.47
10.122 | -94.4708
10.1224 | -94.4723
10.1227 | -94.4737
10.1231 | -94.4743
10.1234 | -94.4749
10.1237 | -94.476
10.1241 | -94.4771
10.1244 | -94.4778
10.1248 | -94.4786
10.1251 | -94.4793
10.1254 | -94.4799
10.1258 | -94.4801
10.1261 | -94.4806
10.1265 | -94.4815
10.1268 | -94.4823
10.1271 | -94.4832
10.1275 | -94.4842
10.1278 | -94.4849
10.1282 | -94.4855
10.1285 | -94.4859
10.1288 | -94.4866
10.1292 | -94.4872
10.1295 | -94.4877
10.1299 | -94.4885
10.1302 | -94.4893
10.1305 | -94.4898
10.1309 | -94.4903
10.1312 | -94.4911
10.1316 | -94.492
10.1319 | -94.4928
10.1322 | -94.4936
10.1326 | -94.4943
10.1329 | -94.4949
10.1333 | -94.4956
10.1336 | -94.4963
10.1339 | -94.4969
10.1343 | -94.4978
10.1346 | -94.4985
10.135 | -94.4995
10.1353 | -94.5001
10.1357 | -94.5008
10.136 | -94.5024
10.1363 | -94.5034
10.1367 | -94.5042
10.137 | -94.5054
10.1374 | -94.5067
10.1377 | -94.5074
10.138 | -94.5082
10.1384 | -94.5089
10.1387 | -94.5098
10.1391 | -94.5107
10.1394 | -94.5114
10.1397 | -94.512
10.1401 | -94.5128
10.1404 | -94.5137
10.1408 | -94.5145
10.1411 | -94.5154
10.1414 | -94.5166
10.1418 | -94.5174
10.1421 | -94.5179
10.1425 | -94.5183
10.1428 | -94.5185
10.1431 | -94.5192
10.1435 | -94.5198
10.1438 | -94.5202
10.1442 | -94.5211
10.1445 | -94.5224
10.1448 | -94.5231
10.1452 | -94.5235
10.1455 | -94.5241
10.1459 | -94.5248
10.1462 | -94.5255
10.1465 | -94.5256
10.1469 | -94.5265
10.1472 | -94.5276
10.1476 | -94.5286
10.1479 | -94.5297
10.1482 | -94.5306
10.1486 | -94.5315
10.1489 | -94.5321
10.1493 | -94.5326
10.1496 | -94.5333
10.15 | -94.534
10.1503 | -94.5347
10.1506 | -94.5352
10.151 | -94.5356
10.1513 | -94.5365
10.1517 | -94.5372
10.152 | -94.5381
10.1523 | -94.5392
10.1527 | -94.5403
10.153 | -94.541
10.1534 | -94.542
10.1537 | -94.5427
10.154 | -94.5436
10.1544 | -94.5444
10.1547 | -94.5454
10.1551 | -94.5463
10.1554 | -94.5474
10.1557 | -94.5484
10.1561 | -94.549
10.1564 | -94.5499
10.1568 | -94.5506
10.1571 | -94.5514
10.1574 | -94.5522
10.1578 | -94.5534
10.1581 | -94.5542
10.1585 | -94.5552
10.1588 | -94.5561
10.1591 | -94.557
10.1595 | -94.5582
10.1598 | -94.5593
10.1602 | -94.5606
10.1605 | -94.5615
10.1608 | -94.5623
10.1612 | -94.5631
10.1615 | -94.564
10.1619 | -94.565
10.1622 | -94.5657
10.1625 | -94.5665
10.1629 | -94.5671
10.1632 | -94.5682
10.1636 | -94.5699
10.1639 | -94.571
10.1643 | -94.5719
10.1646 | -94.5724
10.1649 | -94.5731
10.1653 | -94.5741
10.1656 | -94.5748
10.166 | -94.5756
10.1663 | -94.577
10.1666 | -94.5779
10.167 | -94.5791
10.1673 | -94.5797
10.1677 | -94.5804
10.168 | -94.5811
10.1683 | -94.5827
10.1687 | -94.5852
10.169 | -94.5865
10.1694 | -94.5875
10.1697 | -94.588
10.17 | -94.5884
10.1704 | -94.5889
10.1707 | -94.5895
10.1711 | -94.5902
10.1714 | -94.5909
10.1717 | -94.5921
10.1721 | -94.5932
10.1724 | -94.5943
10.1728 | -94.5951
10.1731 | -94.5958
10.1734 | -94.5967
10.1738 | -94.5979
10.1741 | -94.5987
10.1745 | -94.5994
10.1748 | -94.6003
10.1751 | -94.6018
10.1755 | -94.6031
10.1758 | -94.6042
10.1762 | -94.6051
10.1765 | -94.6053
10.1768 | -94.6061
10.1772 | -94.6071
10.1775 | -94.608
10.1779 | -94.6088
10.1782 | -94.6096
10.1786 | -94.6105
10.1789 | -94.6112
10.1792 | -94.6116
10.1796 | -94.6123
10.1799 | -94.6131
10.1803 | -94.614
10.1806 | -94.6148
10.1809 | -94.6159
10.1813 | -94.617
10.1816 | -94.6178
10.182 | -94.6188
10.1823 | -94.6203
10.1826 | -94.6227
10.183 | -94.6242
10.1833 | -94.6252
10.1837 | -94.6261
10.184 | -94.627
10.1843 | -94.6276
10.1847 | -94.6283
10.185 | -94.6292
10.1854 | -94.6301
10.1857 | -94.6315
10.186 | -94.6325
10.1864 | -94.6332
10.1867 | -94.6341
10.1871 | -94.6352
10.1874 | -94.636
10.1877 | -94.6366
10.1881 | -94.6378
10.1884 | -94.6388
10.1888 | -94.6399
10.1891 | -94.6408
10.1894 | -94.6417
10.1898 | -94.6425
10.1901 | -94.6435
10.1905 | -94.6443
10.1908 | -94.6448
10.1912 | -94.6456
10.1915 | -94.6466
10.1918 | -94.6476
10.1922 | -94.6485
10.1925 | -94.6498
10.1929 | -94.6504
10.1932 | -94.6512
10.1935 | -94.6521
10.1939 | -94.6531
10.1942 | -94.6535
10.1946 | -94.6549
10.1949 | -94.6565
10.1952 | -94.6572
10.1956 | -94.658
10.1959 | -94.6592
10.1963 | -94.6609
10.1966 | -94.6627
10.1969 | -94.6637
10.1973 | -94.6646
10.1976 | -94.6655
10.198 | -94.6662
10.1983 | -94.6672
10.1986 | -94.6685
10.199 | -94.6693
10.1993 | -94.67
10.1997 | -94.671
10.2 | -94.6723
10.2003 | -94.6728
10.2007 | -94.6733
10.201 | -94.6743
10.2014 | -94.6752
10.2017 | -94.6762
10.202 | -94.6771
10.2024 | -94.6779
10.2027 | -94.6785
10.2031 | -94.6793
10.2034 | -94.6804
10.2037 | -94.6816
10.2041 | -94.6825
10.2044 | -94.6833
10.2048 | -94.6842
10.2051 | -94.6849
10.2055 | -94.6858
10.2058 | -94.6864
10.2061 | -94.6868
10.2065 | -94.6873
10.2068 | -94.6883
10.2072 | -94.6891
10.2075 | -94.6908
10.2078 | -94.6922
10.2082 | -94.6932
10.2085 | -94.6945
10.2089 | -94.6956
10.2092 | -94.6967
10.2095 | -94.6979
10.2099 | -94.6987
10.2102 | -94.7
10.2106 | -94.7013
10.2109 | -94.7024
10.2112 | -94.7032
10.2116 | -94.7044
10.2119 | -94.7052
10.2123 | -94.707
10.2126 | -94.7084
10.2129 | -94.7092
10.2133 | -94.71
10.2136 | -94.7106
10.214 | -94.7107
10.2143 | -94.7115
10.2146 | -94.7122
10.215 | -94.7131
10.2153 | -94.714
10.2157 | -94.7154
10.216 | -94.7173
10.2163 | -94.7197
10.2167 | -94.7209
10.217 | -94.7227
10.2174 | -94.7237
10.2177 | -94.7252
10.218 | -94.7264
10.2184 | -94.7268
10.2187 | -94.7277
10.2191 | -94.7284
10.2194 | -94.7282
10.2198 | -94.7283
10.2201 | -94.7285
10.2204 | -94.7285
10.2208 | -94.7286
10.2211 | -94.7293
10.2215 | -94.7304
10.2218 | -94.7315
10.2221 | -94.7325
10.2225 | -94.7337
10.2228 | -94.735
10.2232 | -94.7359
10.2235 | -94.7369
10.2238 | -94.7383
10.2242 | -94.7389
10.2245 | -94.7399
10.2249 | -94.7406
10.2252 | -94.7416
10.2255 | -94.7424
10.2259 | -94.7432
10.2262 | -94.7441
10.2266 | -94.7453
10.2269 | -94.7463
10.2272 | -94.7472
10.2276 | -94.7478
10.2279 | -94.7485
10.2283 | -94.7495
10.2286 | -94.751
10.2289 | -94.7529
10.2293 | -94.7538
10.2296 | -94.7544
10.23 | -94.7552
10.2303 | -94.7561
10.2306 | -94.7569
10.231 | -94.7576
10.2313 | -94.7585
10.2317 | -94.7597
10.232 | -94.7606
10.2323 | -94.7615
10.2327 | -94.7626
10.233 | -94.7635
10.2334 | -94.7646
10.2337 | -94.7657
10.2341 | -94.7672
10.2344 | -94.7692
10.2347 | -94.7702
10.2351 | -94.7711
10.2354 | -94.7719
10.2358 | -94.7729
10.2361 | -94.7739
10.2364 | -94.7746
10.2368 | -94.7754
10.2371 | -94.7757
10.2375 | -94.7763
10.2378 | -94.777
10.2381 | -94.7783
10.2385 | -94.7802
10.2388 | -94.7812
10.2392 | -94.7821
10.2395 | -94.783
10.2398 | -94.7837
10.2402 | -94.7845
10.2405 | -94.785
10.2409 | -94.7856
10.2412 | -94.787
10.2415 | -94.7881
10.2419 | -94.7893
10.2422 | -94.7902
10.2426 | -94.7907
10.2429 | -94.7917
10.2432 | -94.7928
10.2436 | -94.7947
10.2439 | -94.7959
10.2443 | -94.7967
10.2446 | -94.7973
10.2449 | -94.7976
10.2453 | -94.7979
10.2456 | -94.7988
10.246 | -94.7992
10.2463 | -94.7998
10.2467 | -94.8007
10.247 | -94.8018
10.2473 | -94.8026
10.2477 | -94.8034
10.248 | -94.8041
10.2484 | -94.8049
10.2487 | -94.8062
10.249 | -94.8073
10.2494 | -94.8084
10.2497 | -94.8094
10.2501 | -94.8102
10.2504 | -94.8113
10.2507 | -94.8121
10.2511 | -94.8127
10.2514 | -94.8131
10.2518 | -94.8137
10.2521 | -94.8144
10.2524 | -94.8152
10.2528 | -94.8158
10.2531 | -94.8166
10.2535 | -94.8174
10.2538 | -94.8183
10.2541 | -94.8193
10.2545 | -94.8201
10.2548 | -94.8212
10.2552 | -94.822
10.2555 | -94.8225
10.2558 | -94.8236
10.2562 | -94.8242
10.2565 | -94.8247
10.2569 | -94.8259
10.2572 | -94.8268
10.2575 | -94.8277
10.2579 | -94.8286
10.2582 | -94.8296
10.2586 | -94.8309
10.2589 | -94.8322
10.2592 | -94.8335
10.2596 | -94.834
10.2599 | -94.8345
10.2603 | -94.8353
10.2606 | -94.8362
10.2609 | -94.837
10.2613 | -94.8379
10.2616 | -94.8387
10.262 | -94.8396
10.2623 | -94.841
10.2627 | -94.8419
10.263 | -94.8428
10.2633 | -94.8434
10.2637 | -94.8439
10.264 | -94.8447
10.2644 | -94.8457
10.2647 | -94.8464
10.265 | -94.8476
10.2654 | -94.8495
10.2657 | -94.8505
10.2661 | -94.8522
10.2664 | -94.8534
10.2667 | -94.8544
10.2671 | -94.8554
10.2674 | -94.8563
10.2678 | -94.8573
10.2681 | -94.858
10.2684 | -94.8585
10.2688 | -94.8592
10.2691 | -94.8603
10.2695 | -94.862
10.2698 | -94.8635
10.2701 | -94.8649
10.2705 | -94.8657
10.2708 | -94.8665
10.2712 | -94.8679
10.2715 | -94.8688
10.2718 | -94.8692
10.2722 | -94.8697
10.2725 | -94.8708
10.2729 | -94.8716
10.2732 | -94.872
10.2735 | -94.8732
10.2739 | -94.8742
10.2742 | -94.8749
10.2746 | -94.8758
10.2749 | -94.8767
10.2753 | -94.8774
10.2756 | -94.8783
10.2759 | -94.8793
10.2763 | -94.8805
10.2766 | -94.8813
10.277 | -94.8821
10.2773 | -94.8828
10.2776 | -94.8834
10.278 | -94.8846
10.2783 | -94.8859
10.2787 | -94.8864
10.279 | -94.8868
10.2793 | -94.8874
10.2797 | -94.8884
10.28 | -94.8896
10.2804 | -94.8902
10.2807 | -94.8909
10.281 | -94.8922
10.2814 | -94.8929
10.2817 | -94.8936
10.2821 | -94.8944
10.2824 | -94.8953
10.2827 | -94.8964
10.2831 | -94.8978
10.2834 | -94.8984
10.2838 | -94.8992
10.2841 | -94.9
10.2844 | -94.9009
10.2848 | -94.9018
10.2851 | -94.9027
10.2855 | -94.9032
10.2858 | -94.9045
10.2861 | -94.9054
10.2865 | -94.9062
10.2868 | -94.907
10.2872 | -94.9076
10.2875 | -94.9084
10.2878 | -94.9092
10.2882 | -94.91
10.2885 | -94.911
10.2889 | -94.9122
10.2892 | -94.913
10.2896 | -94.9141
10.2899 | -94.9149
10.2902 | -94.9155
10.2906 | -94.9163
10.2909 | -94.9166
10.2913 | -94.9172
10.2916 | -94.9183
10.2919 | -94.919
10.2923 | -94.9199
10.2926 | -94.921
10.293 | -94.9219
10.2933 | -94.9227
10.2936 | -94.9235
10.294 | -94.924
10.2943 | -94.9247
10.2947 | -94.9256
10.295 | -94.9258
10.2953 | -94.9263
10.2957 | -94.9271
10.296 | -94.928
10.2964 | -94.9289
10.2967 | -94.9302
10.297 | -94.9312
10.2974 | -94.932
10.2977 | -94.9329
10.2981 | -94.9337
10.2984 | -94.9344
10.2987 | -94.9352
10.2991 | -94.936
10.2994 | -94.9368
10.2998 | -94.9378
10.3001 | -94.9386
10.3004 | -94.9394
10.3008 | -94.9404
10.3011 | -94.9416
10.3015 | -94.9424
10.3018 | -94.9436
10.3021 | -94.9446
10.3025 | -94.945
10.3028 | -94.9454
10.3032 | -94.9462
10.3035 | -94.9464
10.3039 | -94.9473
10.3042 | -94.948
10.3045 | -94.9488
10.3049 | -94.9497
10.3052 | -94.9506
10.3056 | -94.9514
10.3059 | -94.952
10.3062 | -94.9527
10.3066 | -94.9534
10.3069 | -94.954
10.3073 | -94.9547
10.3076 | -94.9556
10.3079 | -94.9564
10.3083 | -94.9576
10.3086 | -94.959
10.309 | -94.9609
10.3093 | -94.962
10.3096 | -94.963
10.31 | -94.9638
10.3103 | -94.9641
10.3107 | -94.9651
10.311 | -94.9661
10.3113 | -94.9668
10.3117 | -94.9675
10.312 | -94.9676
10.3124 | -94.9681
10.3127 | -94.9694
10.313 | -94.9701
10.3134 | -94.971
10.3137 | -94.9721
10.3141 | -94.9731
10.3144 | -94.9742
10.3147 | -94.9752
10.3151 | -94.9747
10.3154 | -94.9751
10.3158 | -94.9759
10.3161 | -94.9773
10.3164 | -94.9782
10.3168 | -94.9791
10.3171 | -94.9798
10.3175 | -94.9806
10.3178 | -94.9809
10.3182 | -94.9813
10.3185 | -94.9821
10.3188 | -94.9821
10.3192 | -94.9826
10.3195 | -94.9837
10.3199 | -94.9837
10.3202 | -94.9836
10.3205 | -94.9841
10.3209 | -94.9855
10.3212 | -94.9867
10.3216 | -94.9874
10.3219 | -94.9882
10.3222 | -94.989
10.3226 | -94.9901
10.3229 | -94.9911
10.3233 | -94.9923
10.3236 | -94.9932
10.3239 | -94.994
10.3243 | -94.9953
10.3246 | -94.9962
10.325 | -94.9976
10.3253 | -94.9985
10.3256 | -94.9991
10.326 | -94.9995
10.3263 | -95
10.3267 | -95.0006
10.327 | -95.0014
10.3273 | -95.0022
10.3277 | -95.0032
10.328 | -95.0046
10.3284 | -95.0055
10.3287 | -95.0052
10.329 | -95.006
10.3294 | -95.0071
10.3297 | -95.0079
10.3301 | -95.0084
10.3304 | -95.0093
10.3308 | -95.0104
10.3311 | -95.0115
10.3314 | -95.0121
10.3318 | -95.0119
10.3321 | -95.0131
10.3325 | -95.0136
10.3328 | -95.0142
10.3331 | -95.0153
10.3335 | -95.0164
10.3338 | -95.0172
10.3342 | -95.018
10.3345 | -95.0185
10.3348 | -95.019
10.3352 | -95.0182
10.3355 | -95.0183
10.3359 | -95.0195
10.3362 | -95.0209
10.3365 | -95.022
10.3369 | -95.0231
10.3372 | -95.024
10.3376 | -95.0246
10.3379 | -95.0253
10.3382 | -95.0261
10.3386 | -95.0271
10.3389 | -95.0277
10.3393 | -95.0281
10.3396 | -95.029
10.3399 | -95.0298
10.3403 | -95.0306
10.3406 | -95.0314
10.341 | -95.0319
10.3413 | -95.0327
10.3416 | -95.0334
10.342 | -95.034
10.3423 | -95.0343
10.3427 | -95.0348
10.343 | -95.0357
10.3433 | -95.0364
10.3437 | -95.0371
10.344 | -95.0379
10.3444 | -95.0384
10.3447 | -95.0397
10.3451 | -95.0403
10.3454 | -95.0406
10.3457 | -95.0417
10.3461 | -95.0426
10.3464 | -95.0432
10.3468 | -95.0441
10.3471 | -95.0448
10.3474 | -95.0444
10.3478 | -95.0437
10.3481 | -95.0445
10.3485 | -95.0451
10.3488 | -95.0459
10.3491 | -95.0471
10.3495 | -95.0478
10.3498 | -95.0481
10.3502 | -95.0488
10.3505 | -95.0497
10.3508 | -95.0503
10.3512 | -95.0509
10.3515 | -95.0517
10.3519 | -95.0524
10.3522 | -95.0537
10.3525 | -95.0552
10.3529 | -95.0557
10.3532 | -95.0561
10.3536 | -95.0568
10.3539 | -95.0574
10.3542 | -95.0578
10.3546 | -95.0584
10.3549 | -95.0594
10.3553 | -95.0602
10.3556 | -95.0609
10.3559 | -95.0615
10.3563 | -95.0623
10.3566 | -95.0631
10.357 | -95.064
10.3573 | -95.0646
10.3576 | -95.0656
10.358 | -95.0664
10.3583 | -95.0668
10.3587 | -95.0678
10.359 | -95.0685
10.3594 | -95.0693
10.3597 | -95.0698
10.36 | -95.0709
10.3604 | -95.0718
10.3607 | -95.0722
10.3611 | -95.0728
10.3614 | -95.0735
10.3617 | -95.0744
10.3621 | -95.0751
10.3624 | -95.076
10.3628 | -95.0772
10.3631 | -95.0788
10.3634 | -95.0795
10.3638 | -95.0798
10.3641 | -95.0807
10.3645 | -95.0813
10.3648 | -95.082
10.3651 | -95.0828
10.3655 | -95.0834
10.3658 | -95.0837
10.3662 | -95.084
10.3665 | -95.0845
10.3668 | -95.0849
10.3672 | -95.0855
10.3675 | -95.0861
10.3679 | -95.0868
10.3682 | -95.0874
10.3685 | -95.0884
10.3689 | -95.0893
10.3692 | -95.0897
10.3696 | -95.0901
10.3699 | -95.0908
10.3702 | -95.0917
10.3706 | -95.0923
10.3709 | -95.0934
10.3713 | -95.0947
10.3716 | -95.0954
10.3719 | -95.0966
10.3723 | -95.098
10.3726 | -95.0984
10.373 | -95.0992
10.3733 | -95.1014
10.3737 | -95.1023
10.374 | -95.1019
10.3743 | -95.1029
10.3747 | -95.1035
10.375 | -95.104
10.3754 | -95.1047
10.3757 | -95.1054
10.376 | -95.1061
10.3764 | -95.1068
10.3767 | -95.1077
10.3771 | -95.1087
10.3774 | -95.1096
10.3777 | -95.1103
10.3781 | -95.1107
10.3784 | -95.1114
10.3788 | -95.112
10.3791 | -95.1127
10.3794 | -95.1133
10.3798 | -95.1139
10.3801 | -95.1146
10.3805 | -95.1151
10.3808 | -95.1156
10.3811 | -95.1161
10.3815 | -95.117
10.3818 | -95.1178
10.3822 | -95.1183
10.3825 | -95.1188
10.3828 | -95.1195
10.3832 | -95.1202
10.3835 | -95.1206
10.3839 | -95.1208
10.3842 | -95.1221
10.3845 | -95.1231
10.3849 | -95.1236
10.3852 | -95.1246
10.3856 | -95.1252
10.3859 | -95.1257
10.3863 | -95.1263
10.3866 | -95.1273
10.3869 | -95.1283
10.3873 | -95.13
10.3876 | -95.131
10.388 | -95.1315
10.3883 | -95.1322
10.3886 | -95.1327
10.389 | -95.1339
10.3893 | -95.1351
10.3897 | -95.1357
10.39 | -95.1363
10.3903 | -95.1369
10.3907 | -95.1377
10.391 | -95.1385
10.3914 | -95.1391
10.3917 | -95.139
10.392 | -95.1395
10.3924 | -95.1405
10.3927 | -95.1415
10.3931 | -95.1423
10.3934 | -95.1426
10.3937 | -95.1427
10.3941 | -95.1427
10.3944 | -95.1431
10.3948 | -95.1444
10.3951 | -95.1446
10.3954 | -95.1444
10.3958 | -95.1449
10.3961 | -95.1457
10.3965 | -95.1464
10.3968 | -95.1469
10.3971 | -95.1475
10.3975 | -95.1481
10.3978 | -95.1487
10.3982 | -95.1492
10.3985 | -95.1494
10.3988 | -95.15
10.3992 | -95.1504
10.3995 | -95.1509
10.3999 | -95.1512
10.4002 | -95.1514
10.4006 | -95.1521
10.4009 | -95.1528
10.4012 | -95.1534
10.4016 | -95.153
10.4019 | -95.153
10.4023 | -95.1537
10.4026 | -95.1542
10.4029 | -95.1549
10.4033 | -95.1551
10.4036 | -95.1557
10.404 | -95.1562
10.4043 | -95.1568
10.4046 | -95.1574
10.405 | -95.1579
10.4053 | -95.1583
10.4057 | -95.1588
10.406 | -95.1597
10.4063 | -95.1605
10.4067 | -95.1614
10.407 | -95.162
10.4074 | -95.1626
10.4077 | -95.1633
10.408 | -95.1642
10.4084 | -95.1648
10.4087 | -95.1654
10.4091 | -95.1664
10.4094 | -95.1668
10.4097 | -95.1671
10.4101 | -95.1676
10.4104 | -95.1677
10.4108 | -95.1683
10.4111 | -95.1688
10.4114 | -95.1693
10.4118 | -95.1701
10.4121 | -95.1709
10.4125 | -95.1716
10.4128 | -95.1717
10.4131 | -95.1717
10.4135 | -95.1722
10.4138 | -95.1726
10.4142 | -95.1736
10.4145 | -95.1741
10.4149 | -95.1743
10.4152 | -95.1746
10.4155 | -95.1751
10.4159 | -95.176
10.4162 | -95.1766
10.4166 | -95.1769
10.4169 | -95.1775
10.4172 | -95.1774
10.4176 | -95.1774
10.4179 | -95.1778
10.4183 | -95.1783
10.4186 | -95.1787
10.4189 | -95.1798
10.4193 | -95.1822
10.4196 | -95.1842
10.42 | -95.1847
10.4203 | -95.1852
10.4206 | -95.186
10.421 | -95.1861
10.4213 | -95.1866
10.4217 | -95.1869
10.422 | -95.187
10.4223 | -95.1872
10.4227 | -95.1875
10.423 | -95.1882
10.4234 | -95.1889
10.4237 | -95.1895
10.424 | -95.1899
10.4244 | -95.1903
10.4247 | -95.1908
10.4251 | -95.1912
10.4254 | -95.192
10.4257 | -95.1929
10.4261 | -95.194
10.4264 | -95.1959
10.4268 | -95.1969
10.4271 | -95.1972
10.4274 | -95.1978
10.4278 | -95.1978
10.4281 | -95.1985
10.4285 | -95.2
10.4288 | -95.2017
10.4292 | -95.2029
10.4295 | -95.2034
10.4298 | -95.2037
10.4302 | -95.2047
10.4305 | -95.2051
10.4309 | -95.2053
10.4312 | -95.2058
10.4315 | -95.2062
10.4319 | -95.2071
10.4322 | -95.2079
10.4326 | -95.2082
10.4329 | -95.2087
10.4332 | -95.2093
10.4336 | -95.2102
10.4339 | -95.2113
10.4343 | -95.2116
10.4346 | -95.2121
10.4349 | -95.2128
10.4353 | -95.2131
10.4356 | -95.213
10.436 | -95.2129
10.4363 | -95.2132
10.4366 | -95.2137
10.437 | -95.2138
10.4373 | -95.2137
10.4377 | -95.2143
10.438 | -95.2147
10.4383 | -95.2154
10.4387 | -95.2165
10.439 | -95.2168
10.4394 | -95.2173
10.4397 | -95.2181
10.44 | -95.2185
10.4404 | -95.2191
10.4407 | -95.2196
10.4411 | -95.22
10.4414 | -95.2202
10.4417 | -95.2198
10.4421 | -95.2186
10.4424 | -95.2173
10.4428 | -95.2172
10.4431 | -95.2175
10.4435 | -95.2175
10.4438 | -95.2173
10.4441 | -95.2168
10.4445 | -95.2165
10.4448 | -95.2169
10.4452 | -95.2173
10.4455 | -95.2179
10.4458 | -95.2186
10.4462 | -95.2188
10.4465 | -95.2197
10.4469 | -95.2204
10.4472 | -95.2203
10.4475 | -95.2206
10.4479 | -95.2212
10.4482 | -95.2223
10.4486 | -95.2227
10.4489 | -95.2226
10.4492 | -95.2229
10.4496 | -95.2235
10.4499 | -95.2236
10.4503 | -95.2237
10.4506 | -95.224
10.4509 | -95.2241
10.4513 | -95.2236
10.4516 | -95.2234
10.452 | -95.2235
10.4523 | -95.2237
10.4526 | -95.2241
10.453 | -95.2243
10.4533 | -95.2247
10.4537 | -95.2246
10.454 | -95.2247
10.4543 | -95.2254
10.4547 | -95.2261
10.455 | -95.226
10.4554 | -95.2262
10.4557 | -95.2269
10.456 | -95.2274
10.4564 | -95.2286
10.4567 | -95.2298
10.4571 | -95.2302
10.4574 | -95.2307
10.4578 | -95.2313
10.4581 | -95.2316
10.4584 | -95.2323
10.4588 | -95.2329
10.4591 | -95.233
10.4595 | -95.2334
10.4598 | -95.2339
10.4601 | -95.2343
10.4605 | -95.2348
10.4608 | -95.2351
10.4612 | -95.2355
10.4615 | -95.2356
10.4618 | -95.2357
10.4622 | -95.2358
10.4625 | -95.2356
10.4629 | -95.2355
10.4632 | -95.2355
10.4635 | -95.2355
10.4639 | -95.2358
10.4642 | -95.2359
10.4646 | -95.2363
10.4649 | -95.2363
10.4652 | -95.237
10.4656 | -95.2382
10.4659 | -95.239
10.4663 | -95.239
10.4666 | -95.2394
10.4669 | -95.2395
10.4673 | -95.2402
10.4676 | -95.2409
10.468 | -95.2414
10.4683 | -95.2417
10.4686 | -95.2419
10.469 | -95.242
10.4693 | -95.2424
10.4697 | -95.2426
10.47 | -95.2429
10.4703 | -95.2431
10.4707 | -95.2435
10.471 | -95.2438
10.4714 | -95.2439
10.4717 | -95.2427
10.4721 | -95.2416
10.4724 | -95.2418
10.4727 | -95.2417
10.4731 | -95.2416
10.4734 | -95.2418
10.4738 | -95.2418
10.4741 | -95.2419
10.4744 | -95.2419
10.4748 | -95.2419
10.4751 | -95.2425
10.4755 | -95.2423
10.4758 | -95.242
10.4761 | -95.2422
10.4765 | -95.2426
10.4768 | -95.2434
10.4772 | -95.2442
10.4775 | -95.2447
10.4778 | -95.245
10.4782 | -95.2451
10.4785 | -95.2452
10.4789 | -95.2453
10.4792 | -95.2454
10.4795 | -95.2455
10.4799 | -95.2456
10.4802 | -95.2459
10.4806 | -95.2464
10.4809 | -95.2468
10.4812 | -95.2474
10.4816 | -95.248
10.4819 | -95.2477
10.4823 | -95.2481
10.4826 | -95.248
10.4829 | -95.2481
10.4833 | -95.2495
10.4836 | -95.2504
10.484 | -95.2502
10.4843 | -95.2499
10.4847 | -95.2498
10.485 | -95.2501
10.4853 | -95.2502
10.4857 | -95.2501
10.486 | -95.2505
10.4864 | -95.2508
10.4867 | -95.2507
10.487 | -95.2507
10.4874 | -95.2508
10.4877 | -95.2515
10.4881 | -95.2532
10.4884 | -95.2544
10.4887 | -95.2544
10.4891 | -95.2548
10.4894 | -95.2551
10.4898 | -95.2559
10.4901 | -95.2573
10.4904 | -95.2576
10.4908 | -95.2577
10.4911 | -95.2579
10.4915 | -95.258
10.4918 | -95.258
10.4921 | -95.2576
10.4925 | -95.2576
10.4928 | -95.258
10.4932 | -95.2582
10.4935 | -95.2589
10.4938 | -95.2597
10.4942 | -95.2599
10.4945 | -95.2597
10.4949 | -95.2589
10.4952 | -95.2584
10.4955 | -95.2587
10.4959 | -95.259
10.4962 | -95.2594
10.4966 | -95.2601
10.4969 | -95.2607
10.4972 | -95.2608
10.4976 | -95.2613
10.4979 | -95.2615
10.4983 | -95.2609
10.4986 | -95.261
10.499 | -95.2609
10.4993 | -95.2606
10.4996 | -95.2603
10.5 | -95.2604
10.5003 | -95.2611
10.5007 | -95.2617
10.501 | -95.2618
10.5013 | -95.2617
10.5017 | -95.2615
10.502 | -95.2616
10.5024 | -95.2622
10.5027 | -95.2624
10.503 | -95.2623
10.5034 | -95.2619
10.5037 | -95.2617
10.5041 | -95.2618
10.5044 | -95.2621
10.5047 | -95.262
10.5051 | -95.2621
10.5054 | -95.262
10.5058 | -95.2618
10.5061 | -95.2615
10.5064 | -95.2617
10.5068 | -95.2626
10.5071 | -95.2634
10.5075 | -95.2633
10.5078 | -95.2636
10.5081 | -95.2636
10.5085 | -95.2636
10.5088 | -95.2633
10.5092 | -95.2631
10.5095 | -95.263
10.5098 | -95.2631
10.5102 | -95.2631
10.5105 | -95.2631
10.5109 | -95.2628
10.5112 | -95.2625
10.5115 | -95.2619
10.5119 | -95.2617
10.5122 | -95.2616
10.5126 | -95.2609
10.5129 | -95.2597
10.5133 | -95.2589
10.5136 | -95.2592
10.5139 | -95.2594
10.5143 | -95.2594
10.5146 | -95.2594
10.515 | -95.2592
10.5153 | -95.2592
10.5156 | -95.2591
10.516 | -95.2591
10.5163 | -95.2588
10.5167 | -95.2591
10.517 | -95.2595
10.5173 | -95.26
10.5177 | -95.26
10.518 | -95.2596
10.5184 | -95.2584
10.5187 | -95.258
10.519 | -95.2579
10.5194 | -95.2579
10.5197 | -95.2581
10.5201 | -95.2584
10.5204 | -95.2585
10.5207 | -95.2583
10.5211 | -95.2565
10.5214 | -95.2563
10.5218 | -95.2562
10.5221 | -95.256
10.5224 | -95.2556
10.5228 | -95.2551
10.5231 | -95.2546
10.5235 | -95.2544
10.5238 | -95.2535
10.5241 | -95.2517
10.5245 | -95.2502
10.5248 | -95.2495
10.5252 | -95.2489
10.5255 | -95.2482
10.5259 | -95.2477
10.5262 | -95.2476
10.5265 | -95.2473
10.5269 | -95.2467
10.5272 | -95.2462
10.5276 | -95.246
10.5279 | -95.2459
10.5282 | -95.2458
10.5286 | -95.2454
10.5289 | -95.2451
10.5293 | -95.2453
10.5296 | -95.246
10.5299 | -95.2461
10.5303 | -95.246
10.5306 | -95.2466
10.531 | -95.246
10.5313 | -95.247
10.5316 | -95.2476
10.532 | -95.2471
10.5323 | -95.2464
10.5327 | -95.2462
10.533 | -95.246
10.5333 | -95.2458
10.5337 | -95.2457
10.534 | -95.2456
10.5344 | -95.2454
10.5347 | -95.2452
10.535 | -95.2452
10.5354 | -95.2447
10.5357 | -95.2446
10.5361 | -95.2445
10.5364 | -95.2442
10.5367 | -95.2441
10.5371 | -95.2438
10.5374 | -95.2434
10.5378 | -95.2432
10.5381 | -95.243
10.5384 | -95.243
10.5388 | -95.2428
10.5391 | -95.242
10.5395 | -95.242
10.5398 | -95.2424
10.5402 | -95.2425
10.5405 | -95.2419
10.5408 | -95.2411
10.5412 | -95.2407
10.5415 | -95.2408
10.5419 | -95.2404
10.5422 | -95.2403
10.5425 | -95.2403
10.5429 | -95.2405
10.5432 | -95.2406
10.5436 | -95.2403
10.5439 | -95.2395
10.5442 | -95.2393
10.5446 | -95.2395
10.5449 | -95.2403
10.5453 | -95.2404
10.5456 | -95.2404
10.5459 | -95.2419
10.5463 | -95.2421
10.5466 | -95.2416
10.547 | -95.2409
10.5473 | -95.24
10.5476 | -95.2399
10.548 | -95.24
10.5483 | -95.24
10.5487 | -95.2395
10.549 | -95.2389
10.5493 | -95.2386
10.5497 | -95.238
10.55 | -95.237
10.5504 | -95.2365
10.5507 | -95.2364
10.551 | -95.2358
10.5514 | -95.2351
10.5517 | -95.2349
10.5521 | -95.2346
10.5524 | -95.2346
10.5527 | -95.2347
10.5531 | -95.2346
10.5534 | -95.236
10.5538 | -95.2372
10.5541 | -95.2377
10.5545 | -95.2376
10.5548 | -95.2372
10.5551 | -95.2369
10.5555 | -95.2366
10.5558 | -95.2364
10.5562 | -95.2361
10.5565 | -95.2357
10.5568 | -95.2352
10.5572 | -95.2352
10.5575 | -95.2349
10.5579 | -95.2351
10.5582 | -95.2351
10.5585 | -95.2351
10.5589 | -95.2348
10.5592 | -95.2346
10.5596 | -95.2344
10.5599 | -95.2339
10.5602 | -95.2337
10.5606 | -95.2334
10.5609 | -95.2331
10.5613 | -95.2329
10.5616 | -95.2326
10.5619 | -95.2323
10.5623 | -95.2319
10.5626 | -95.2313
10.563 | -95.2306
10.5633 | -95.2298
10.5636 | -95.2297
10.564 | -95.23
10.5643 | -95.2297
10.5647 | -95.2294
10.565 | -95.229
10.5653 | -95.2285
10.5657 | -95.2278
10.566 | -95.2277
10.5664 | -95.2277
10.5667 | -95.2272
10.567 | -95.2269
10.5674 | -95.2268
10.5677 | -95.2268
10.5681 | -95.2267
10.5684 | -95.2263
10.5688 | -95.2261
10.5691 | -95.2253
10.5694 | -95.2245
10.5698 | -95.2237
10.5701 | -95.2239
10.5705 | -95.2235
10.5708 | -95.2233
10.5711 | -95.2229
10.5715 | -95.2226
10.5718 | -95.2223
10.5722 | -95.2225
10.5725 | -95.2227
10.5728 | -95.2221
10.5732 | -95.2213
10.5735 | -95.2203
10.5739 | -95.2197
10.5742 | -95.2191
10.5745 | -95.2187
10.5749 | -95.2181
10.5752 | -95.2174
10.5756 | -95.2168
10.5759 | -95.2163
10.5762 | -95.2156
10.5766 | -95.2155
10.5769 | -95.2157
10.5773 | -95.2154
10.5776 | -95.2149
10.5779 | -95.2138
10.5783 | -95.2125
10.5786 | -95.2114
10.579 | -95.2106
10.5793 | -95.2099
10.5796 | -95.2091
10.58 | -95.2085
10.5803 | -95.2081
10.5807 | -95.2077
10.581 | -95.2072
10.5813 | -95.2066
10.5817 | -95.2065
10.582 | -95.2068
10.5824 | -95.2065
10.5827 | -95.2061
10.5831 | -95.206
10.5834 | -95.2055
10.5837 | -95.205
10.5841 | -95.2045
10.5844 | -95.2037
10.5848 | -95.2027
10.5851 | -95.2023
10.5854 | -95.2025
10.5858 | -95.2025
10.5861 | -95.2022
10.5865 | -95.2022
10.5868 | -95.2017
10.5871 | -95.2013
10.5875 | -95.2007
10.5878 | -95.2002
10.5882 | -95.1997
10.5885 | -95.1995
10.5888 | -95.1997
10.5892 | -95.1994
10.5895 | -95.1993
10.5899 | -95.1987
10.5902 | -95.1977
10.5905 | -95.1969
10.5909 | -95.1967
10.5912 | -95.1966
10.5916 | -95.1961
10.5919 | -95.1955
10.5922 | -95.1945
10.5926 | -95.1938
10.5929 | -95.1934
10.5933 | -95.1929
10.5936 | -95.1927
10.5939 | -95.1924
10.5943 | -95.1915
10.5946 | -95.1903
10.595 | -95.19
10.5953 | -95.19
10.5956 | -95.189
10.596 | -95.1879
10.5963 | -95.1872
10.5967 | -95.1868
10.597 | -95.1865
10.5974 | -95.1864
10.5977 | -95.186
10.598 | -95.1857
10.5984 | -95.1853
10.5987 | -95.1848
10.5991 | -95.1843
10.5994 | -95.1835
10.5997 | -95.1829
10.6001 | -95.1827
10.6004 | -95.1828
10.6008 | -95.1828
10.6011 | -95.1823
10.6014 | -95.1814
10.6018 | -95.1809
10.6021 | -95.1805
10.6025 | -95.1803
10.6028 | -95.1797
10.6031 | -95.1793
10.6035 | -95.1787
10.6038 | -95.1784
10.6042 | -95.1781
10.6045 | -95.1779
10.6048 | -95.1776
10.6052 | -95.1768
10.6055 | -95.1758
10.6059 | -95.1753
10.6062 | -95.1751
10.6065 | -95.1744
10.6069 | -95.1735
10.6072 | -95.173
10.6076 | -95.1726
10.6079 | -95.1722
10.6082 | -95.1715
10.6086 | -95.1712
10.6089 | -95.1708
10.6093 | -95.1703
10.6096 | -95.1679
10.6099 | -95.1671
10.6103 | -95.1667
10.6106 | -95.1662
10.611 | -95.1656
10.6113 | -95.1651
10.6117 | -95.1648
10.612 | -95.1642
10.6123 | -95.1635
10.6127 | -95.1632
10.613 | -95.1632
10.6134 | -95.1621
10.6137 | -95.1608
10.614 | -95.1603
10.6144 | -95.1595
10.6147 | -95.1591
10.6151 | -95.1583
10.6154 | -95.1576
10.6157 | -95.1572
10.6161 | -95.1566
10.6164 | -95.1556
10.6168 | -95.155
10.6171 | -95.1547
10.6174 | -95.1542
10.6178 | -95.1536
10.6181 | -95.1536
10.6185 | -95.1536
10.6188 | -95.1534
10.6191 | -95.1528
10.6195 | -95.1526
10.6198 | -95.1516
10.6202 | -95.1509
10.6205 | -95.1505
10.6208 | -95.1496
10.6212 | -95.1489
10.6215 | -95.1481
10.6219 | -95.1477
10.6222 | -95.1473
10.6225 | -95.1468
10.6229 | -95.1465
10.6232 | -95.1462
10.6236 | -95.1452
10.6239 | -95.144
10.6243 | -95.1432
10.6246 | -95.1427
10.6249 | -95.1422
10.6253 | -95.1418
10.6256 | -95.1408
10.626 | -95.1401
10.6263 | -95.1393
10.6266 | -95.1384
10.627 | -95.1378
10.6273 | -95.1369
10.6277 | -95.1354
10.628 | -95.1344
10.6283 | -95.134
10.6287 | -95.134
10.629 | -95.1331
10.6294 | -95.1324
10.6297 | -95.1318
10.63 | -95.1311
10.6304 | -95.1304
10.6307 | -95.1295
10.6311 | -95.1287
10.6314 | -95.1283
10.6317 | -95.1278
10.6321 | -95.1277
10.6324 | -95.1272
10.6328 | -95.1272
10.6331 | -95.1269
10.6334 | -95.1263
10.6338 | -95.1256
10.6341 | -95.1252
10.6345 | -95.1247
10.6348 | -95.1243
10.6351 | -95.1239
10.6355 | -95.1237
10.6358 | -95.1239
10.6362 | -95.1235
10.6365 | -95.123
10.6368 | -95.1225
10.6372 | -95.1219
10.6375 | -95.1214
10.6379 | -95.1209
10.6382 | -95.1204
10.6386 | -95.1201
10.6389 | -95.1197
10.6392 | -95.1192
10.6396 | -95.1186
10.6399 | -95.1181
10.6403 | -95.1176
10.6406 | -95.1172
10.6409 | -95.1168
10.6413 | -95.1161
10.6416 | -95.1156
10.642 | -95.1147
10.6423 | -95.1137
10.6426 | -95.1132
10.643 | -95.1125
10.6433 | -95.1118
10.6437 | -95.111
10.644 | -95.1105
10.6443 | -95.1107
10.6447 | -95.1113
10.645 | -95.1108
10.6454 | -95.1095
10.6457 | -95.109
10.646 | -95.1083
10.6464 | -95.108
10.6467 | -95.1074
10.6471 | -95.1072
10.6474 | -95.1066
10.6477 | -95.1061
10.6481 | -95.106
10.6484 | -95.1056
10.6488 | -95.1051
10.6491 | -95.1048
10.6494 | -95.1045
10.6498 | -95.1039
10.6501 | -95.1037
10.6505 | -95.1032
10.6508 | -95.1029
10.6511 | -95.1018
10.6515 | -95.1013
10.6518 | -95.101
10.6522 | -95.1005
10.6525 | -95.0998
10.6529 | -95.0992
10.6532 | -95.0985
10.6535 | -95.0979
10.6539 | -95.0974
10.6542 | -95.0969
10.6546 | -95.0961
10.6549 | -95.0954
10.6552 | -95.0952
10.6556 | -95.0944
10.6559 | -95.0952
10.6563 | -95.0952
10.6566 | -95.0946
10.6569 | -95.094
10.6573 | -95.0936
10.6576 | -95.0933
10.658 | -95.0927
10.6583 | -95.0923
10.6586 | -95.091
10.659 | -95.0911
10.6593 | -95.091
10.6597 | -95.0902
10.66 | -95.0899
10.6603 | -95.0895
10.6607 | -95.0887
10.661 | -95.0882
10.6614 | -95.0877
10.6617 | -95.0871
10.662 | -95.0865
10.6624 | -95.0858
10.6627 | -95.0853
10.6631 | -95.0853
10.6634 | -95.0854
10.6637 | -95.0849
10.6641 | -95.0844
10.6644 | -95.0838
10.6648 | -95.0831
10.6651 | -95.0824
10.6654 | -95.0818
10.6658 | -95.0812
10.6661 | -95.081
10.6665 | -95.0806
10.6668 | -95.0802
10.6672 | -95.0801
10.6675 | -95.0801
10.6678 | -95.0795
10.6682 | -95.0788
10.6685 | -95.0781
10.6689 | -95.0773
10.6692 | -95.0768
10.6695 | -95.0763
10.6699 | -95.0758
10.6702 | -95.0756
10.6706 | -95.0749
10.6709 | -95.0743
10.6712 | -95.074
10.6716 | -95.0733
10.6719 | -95.0728
10.6723 | -95.0723
10.6726 | -95.0715
10.6729 | -95.071
10.6733 | -95.0708
10.6736 | -95.0705
10.674 | -95.07
10.6743 | -95.0699
10.6746 | -95.0692
10.675 | -95.0683
10.6753 | -95.0678
10.6757 | -95.0673
10.676 | -95.0672
10.6763 | -95.0669
10.6767 | -95.0666
10.677 | -95.0661
10.6774 | -95.0655
10.6777 | -95.0651
10.678 | -95.0651
10.6784 | -95.0647
10.6787 | -95.0643
10.6791 | -95.0634
10.6794 | -95.0628
10.6798 | -95.0622
10.6801 | -95.0614
10.6804 | -95.0598
10.6808 | -95.0583
10.6811 | -95.058
10.6815 | -95.0574
10.6818 | -95.0569
10.6821 | -95.0564
10.6825 | -95.056
10.6828 | -95.0558
10.6832 | -95.0554
10.6835 | -95.0547
10.6838 | -95.0537
10.6842 | -95.0532
10.6845 | -95.053
10.6849 | -95.0528
10.6852 | -95.0525
10.6855 | -95.052
10.6859 | -95.0512
10.6862 | -95.0505
10.6866 | -95.0501
10.6869 | -95.0499
10.6872 | -95.0496
10.6876 | -95.0487
10.6879 | -95.0478
10.6883 | -95.0478
10.6886 | -95.0474
10.6889 | -95.0468
10.6893 | -95.047
10.6896 | -95.0473
10.69 | -95.047
10.6903 | -95.0465
10.6906 | -95.0459
10.691 | -95.0452
10.6913 | -95.0448
10.6917 | -95.0442
10.692 | -95.0436
10.6923 | -95.0432
10.6927 | -95.043
10.693 | -95.0424
10.6934 | -95.0421
10.6937 | -95.0415
10.6941 | -95.0405
10.6944 | -95.0393
10.6947 | -95.0386
10.6951 | -95.0382
10.6954 | -95.038
10.6958 | -95.0376
10.6961 | -95.0371
10.6964 | -95.0364
10.6968 | -95.036
10.6971 | -95.0358
10.6975 | -95.0358
10.6978 | -95.0353
10.6981 | -95.0348
10.6985 | -95.0346
10.6988 | -95.0341
10.6992 | -95.0338
10.6995 | -95.0333
10.6998 | -95.0337
10.7002 | -95.0333
10.7005 | -95.0328
10.7009 | -95.032
10.7012 | -95.0313
10.7015 | -95.0309
10.7019 | -95.0304
10.7022 | -95.03
10.7026 | -95.0297
10.7029 | -95.0293
10.7032 | -95.0286
10.7036 | -95.0281
10.7039 | -95.0276
10.7043 | -95.0271
10.7046 | -95.0267
10.7049 | -95.0263
10.7053 | -95.0256
10.7056 | -95.0251
10.706 | -95.0246
10.7063 | -95.0238
10.7066 | -95.0234
10.707 | -95.0231
10.7073 | -95.0231
10.7077 | -95.0233
10.708 | -95.0232
10.7084 | -95.0232
10.7087 | -95.0228
10.709 | -95.0229
10.7094 | -95.0228
10.7097 | -95.0225
10.7101 | -95.022
10.7104 | -95.0219
10.7107 | -95.0218
10.7111 | -95.0215
10.7114 | -95.0206
10.7118 | -95.02
10.7121 | -95.0194
10.7124 | -95.0187
10.7128 | -95.0179
10.7131 | -95.0173
10.7135 | -95.0172
10.7138 | -95.017
10.7141 | -95.0168
10.7145 | -95.0164
10.7148 | -95.0162
10.7152 | -95.0161
10.7155 | -95.0156
10.7158 | -95.0148
10.7162 | -95.0139
10.7165 | -95.0136
10.7169 | -95.0134
10.7172 | -95.0129
10.7175 | -95.0124
10.7179 | -95.0121
10.7182 | -95.0116
10.7186 | -95.0109
10.7189 | -95.0103
10.7192 | -95.0097
10.7196 | -95.009
10.7199 | -95.009
10.7203 | -95.0094
10.7206 | -95.009
10.721 | -95.0095
10.7213 | -95.0102
10.7216 | -95.0093
10.722 | -95.0089
10.7223 | -95.0095
10.7227 | -95.01
10.723 | -95.0101
10.7233 | -95.0096
10.7237 | -95.0096
10.724 | -95.0095
10.7244 | -95.0091
10.7247 | -95.0087
10.725 | -95.008
10.7254 | -95.0076
10.7257 | -95.0073
10.7261 | -95.0065
10.7264 | -95.0057
10.7267 | -95.0052
10.7271 | -95.0048
10.7274 | -95.0037
10.7278 | -95.0032
10.7281 | -95.0028
10.7284 | -95.0024
10.7288 | -95.002
10.7291 | -95.0015
10.7295 | -95.0012
10.7298 | -95.0009
10.7301 | -95.0005
10.7305 | -95.0002
10.7308 | -94.9998
10.7312 | -94.9993
10.7315 | -94.9989
10.7318 | -94.9982
10.7322 | -94.9978
10.7325 | -94.9973
10.7329 | -94.9968
10.7332 | -94.9964
10.7335 | -94.9959
10.7339 | -94.9953
10.7342 | -94.9949
10.7346 | -94.9945
10.7349 | -94.994
10.7353 | -94.9933
10.7356 | -94.9924
10.7359 | -94.9919
10.7363 | -94.9915
10.7366 | -94.9913
10.737 | -94.9912
10.7373 | -94.9915
10.7376 | -94.9912
10.738 | -94.9912
10.7383 | -94.9912
10.7387 | -94.9908
10.739 | -94.9904
10.7393 | -94.99
10.7397 | -94.9898
10.74 | -94.9895
10.7404 | -94.989
10.7407 | -94.9887
10.741 | -94.9889
10.7414 | -94.9889
10.7417 | -94.9884
10.7421 | -94.9881
10.7424 | -94.9877
10.7427 | -94.9871
10.7431 | -94.9872
10.7434 | -94.9865
10.7438 | -94.9857
10.7441 | -94.9858
10.7444 | -94.9853
10.7448 | -94.9847
10.7451 | -94.9843
10.7455 | -94.9839
10.7458 | -94.9834
10.7461 | -94.9829
10.7465 | -94.9824
10.7468 | -94.9819
10.7472 | -94.9816
10.7475 | -94.9813
10.7478 | -94.9809
10.7482 | -94.9804
10.7485 | -94.9807
10.7489 | -94.9816
10.7492 | -94.9821
10.7496 | -94.9821
10.7499 | -94.9821
10.7502 | -94.982
10.7506 | -94.9816
10.7509 | -94.9809
10.7513 | -94.9806
10.7516 | -94.981
10.7519 | -94.9808
10.7523 | -94.9805
10.7526 | -94.9802
10.753 | -94.9799
10.7533 | -94.9795
10.7536 | -94.9789
10.754 | -94.9784
10.7543 | -94.9782
10.7547 | -94.9778
10.755 | -94.9775
10.7553 | -94.9774
10.7557 | -94.9771
10.756 | -94.9768
10.7564 | -94.9766
10.7567 | -94.9762
10.757 | -94.9759
10.7574 | -94.9751
10.7577 | -94.9749
10.7581 | -94.9745
10.7584 | -94.9742
10.7587 | -94.9739
10.7591 | -94.9736
10.7594 | -94.9731
10.7598 | -94.973
10.7601 | -94.9727
10.7604 | -94.9723
10.7608 | -94.9721
10.7611 | -94.9717
10.7615 | -94.9715
10.7618 | -94.9718
10.7621 | -94.9722
10.7625 | -94.9723
10.7628 | -94.9724
10.7632 | -94.9725
10.7635 | -94.9719
10.7639 | -94.9718
10.7642 | -94.9729
10.7645 | -94.9734
10.7649 | -94.9733
10.7652 | -94.9733
10.7656 | -94.9736
10.7659 | -94.9734
10.7662 | -94.9736
10.7666 | -94.9735
10.7669 | -94.974
10.7673 | -94.9738
10.7676 | -94.9736
10.7679 | -94.973
10.7683 | -94.9725
10.7686 | -94.9725
10.769 | -94.9725
10.7693 | -94.9728
10.7696 | -94.9728
10.77 | -94.9727
10.7703 | -94.9722
10.7707 | -94.9723
10.771 | -94.9718
10.7713 | -94.9718
10.7717 | -94.9717
10.772 | -94.9712
10.7724 | -94.9709
10.7727 | -94.9705
10.773 | -94.9704
10.7734 | -94.9707
10.7737 | -94.9708
10.7741 | -94.9707
10.7744 | -94.9703
10.7747 | -94.9702
10.7751 | -94.9695
10.7754 | -94.9689
10.7758 | -94.9688
10.7761 | -94.9686
10.7764 | -94.9682
10.7768 | -94.9687
10.7771 | -94.9683
10.7775 | -94.9675
10.7778 | -94.9674
10.7782 | -94.9677
10.7785 | -94.9678
10.7788 | -94.9675
10.7792 | -94.967
10.7795 | -94.9669
10.7799 | -94.9667
10.7802 | -94.9661
10.7805 | -94.9662
10.7809 | -94.9664
10.7812 | -94.967
10.7816 | -94.9672
10.7819 | -94.9672
10.7822 | -94.9669
10.7826 | -94.9665
10.7829 | -94.9663
10.7833 | -94.9663
10.7836 | -94.9665
10.7839 | -94.9661
10.7843 | -94.9661
10.7846 | -94.9663
10.785 | -94.9664
10.7853 | -94.9664
10.7856 | -94.9665
10.786 | -94.9664
10.7863 | -94.9667
10.7867 | -94.9668
10.787 | -94.9671
10.7873 | -94.9668
10.7877 | -94.9667
10.788 | -94.9663
10.7884 | -94.966
10.7887 | -94.9658
10.789 | -94.9655
10.7894 | -94.9651
10.7897 | -94.9648
10.7901 | -94.9646
10.7904 | -94.965
10.7907 | -94.965
10.7911 | -94.9649
10.7914 | -94.9648
10.7918 | -94.9652
10.7921 | -94.9652
10.7925 | -94.965
10.7928 | -94.965
10.7931 | -94.965
10.7935 | -94.9649
10.7938 | -94.965
10.7942 | -94.9653
10.7945 | -94.9652
10.7948 | -94.9651
10.7952 | -94.9651
10.7955 | -94.9662
10.7959 | -94.9666
10.7962 | -94.9671
10.7965 | -94.9669
10.7969 | -94.9668
10.7972 | -94.9662
10.7976 | -94.966
10.7979 | -94.9657
10.7982 | -94.9656
10.7986 | -94.9656
10.7989 | -94.9655
10.7993 | -94.9656
10.7996 | -94.9661
10.7999 | -94.9665
10.8003 | -94.9665
10.8006 | -94.9665
10.801 | -94.9665
10.8013 | -94.9664
10.8016 | -94.9663
10.802 | -94.9663
10.8023 | -94.9664
10.8027 | -94.9674
10.803 | -94.9677
10.8033 | -94.9675
10.8037 | -94.9673
10.804 | -94.9673
10.8044 | -94.967
10.8047 | -94.9669
10.805 | -94.9667
10.8054 | -94.9676
10.8057 | -94.968
10.8061 | -94.9687
10.8064 | -94.9688
10.8068 | -94.9687
10.8071 | -94.9687
10.8074 | -94.9693
10.8078 | -94.9695
10.8081 | -94.9694
10.8085 | -94.9694
10.8088 | -94.9695
10.8091 | -94.9693
10.8095 | -94.9687
10.8098 | -94.9688
10.8102 | -94.9695
10.8105 | -94.9699
10.8108 | -94.9698
10.8112 | -94.9697
10.8115 | -94.9703
10.8119 | -94.9706
10.8122 | -94.9705
10.8125 | -94.9705
10.8129 | -94.9706
10.8132 | -94.9707
10.8136 | -94.9709
10.8139 | -94.9709
10.8142 | -94.9708
10.8146 | -94.9709
10.8149 | -94.9709
10.8153 | -94.9715
10.8156 | -94.9715
10.8159 | -94.9714
10.8163 | -94.972
10.8166 | -94.9727
10.817 | -94.9727
10.8173 | -94.9727
10.8176 | -94.9728
10.818 | -94.9728
10.8183 | -94.9728
10.8187 | -94.9728
10.819 | -94.9727
10.8194 | -94.9726
10.8197 | -94.972
10.82 | -94.9717
10.8204 | -94.9721
10.8207 | -94.9722
10.8211 | -94.9723
10.8214 | -94.9725
10.8217 | -94.9728
10.8221 | -94.9731
10.8224 | -94.9731
10.8228 | -94.9736
10.8231 | -94.974
10.8234 | -94.9739
10.8238 | -94.9741
10.8241 | -94.9739
10.8245 | -94.9738
10.8248 | -94.9739
10.8251 | -94.974
10.8255 | -94.9742
10.8258 | -94.975
10.8262 | -94.9752
10.8265 | -94.9753
10.8268 | -94.9755
10.8272 | -94.9759
10.8275 | -94.976
10.8279 | -94.9762
10.8282 | -94.9762
10.8285 | -94.9758
10.8289 | -94.976
10.8292 | -94.9762
10.8296 | -94.9761
10.8299 | -94.9761
10.8302 | -94.9763
10.8306 | -94.9764
10.8309 | -94.9765
10.8313 | -94.977
10.8316 | -94.9772
10.8319 | -94.9777
10.8323 | -94.9781
10.8326 | -94.9781
10.833 | -94.978
10.8333 | -94.9778
10.8337 | -94.9777
10.834 | -94.9779
10.8343 | -94.978
10.8347 | -94.9781
10.835 | -94.9783
10.8354 | -94.9786
10.8357 | -94.9791
10.836 | -94.9795
10.8364 | -94.9799
10.8367 | -94.9801
10.8371 | -94.9803
10.8374 | -94.9805
10.8377 | -94.9807
10.8381 | -94.9809
10.8384 | -94.9812
10.8388 | -94.9812
10.8391 | -94.9814
10.8394 | -94.9815
10.8398 | -94.9819
10.8401 | -94.9837
10.8405 | -94.9848
10.8408 | -94.9852
10.8411 | -94.9852
10.8415 | -94.9852
10.8418 | -94.9852
10.8422 | -94.9854
10.8425 | -94.9857
10.8428 | -94.9863
10.8432 | -94.9866
10.8435 | -94.9869
10.8439 | -94.987
10.8442 | -94.9873
10.8445 | -94.9878
10.8449 | -94.9883
10.8452 | -94.9887
10.8456 | -94.9895
10.8459 | -94.9897
10.8462 | -94.9896
10.8466 | -94.989
10.8469 | -94.9894
10.8473 | -94.99
10.8476 | -94.9901
10.848 | -94.9901
10.8483 | -94.9903
10.8486 | -94.9908
10.849 | -94.9912
10.8493 | -94.9915
10.8497 | -94.9919
10.85 | -94.992
10.8503 | -94.9922
10.8507 | -94.9925
10.851 | -94.9928
10.8514 | -94.9931
10.8517 | -94.9934
10.852 | -94.9935
10.8524 | -94.9938
10.8527 | -94.9941
10.8531 | -94.9945
10.8534 | -94.9953
10.8537 | -94.9959
10.8541 | -94.996
10.8544 | -94.996
10.8548 | -94.9964
10.8551 | -94.9966
10.8554 | -94.9966
10.8558 | -94.9968
10.8561 | -94.997
10.8565 | -94.9967
10.8568 | -94.9971
10.8571 | -94.9976
10.8575 | -94.9984
10.8578 | -94.9995
10.8582 | -94.9999
10.8585 | -95.0003
10.8588 | -95.0008
10.8592 | -95.0012
10.8595 | -95.0013
10.8599 | -95.0014
10.8602 | -95.0017
10.8605 | -95.0019
10.8609 | -95.0022
10.8612 | -95.0026
10.8616 | -95.0027
10.8619 | -95.0032
10.8623 | -95.0038
10.8626 | -95.0043
10.8629 | -95.0048
10.8633 | -95.0052
10.8636 | -95.0054
10.864 | -95.0052
10.8643 | -95.0054
10.8646 | -95.0056
10.865 | -95.006
10.8653 | -95.0063
10.8657 | -95.0066
10.866 | -95.007
10.8663 | -95.0076
10.8667 | -95.0082
10.867 | -95.0088
10.8674 | -95.0092
10.8677 | -95.0097
10.868 | -95.0101
10.8684 | -95.0106
10.8687 | -95.0112
10.8691 | -95.0113
10.8694 | -95.011
10.8697 | -95.011
10.8701 | -95.0112
10.8704 | -95.0111
10.8708 | -95.0119
10.8711 | -95.0128
10.8714 | -95.0135
10.8718 | -95.0145
10.8721 | -95.015
10.8725 | -95.0154
10.8728 | -95.0161
10.8731 | -95.0164
10.8735 | -95.0168
10.8738 | -95.0175
10.8742 | -95.0185
10.8745 | -95.0189
10.8749 | -95.019
10.8752 | -95.0193
10.8755 | -95.0199
10.8759 | -95.0211
10.8762 | -95.0213
10.8766 | -95.0215
10.8769 | -95.0219
10.8772 | -95.0222
10.8776 | -95.0223
10.8779 | -95.0225
10.8783 | -95.0228
10.8786 | -95.0233
10.8789 | -95.0237
10.8793 | -95.024
10.8796 | -95.0243
10.88 | -95.0246
10.8803 | -95.0254
10.8806 | -95.0265
10.881 | -95.0271
10.8813 | -95.0277
10.8817 | -95.0279
10.882 | -95.0276
10.8823 | -95.027
10.8827 | -95.0272
10.883 | -95.0276
10.8834 | -95.028
10.8837 | -95.0283
10.884 | -95.0282
10.8844 | -95.0281
10.8847 | -95.0285
10.8851 | -95.0284
10.8854 | -95.0281
10.8857 | -95.0284
10.8861 | -95.029
10.8864 | -95.0295
10.8868 | -95.0299
10.8871 | -95.0303
10.8874 | -95.0304
10.8878 | -95.0305
10.8881 | -95.0306
10.8885 | -95.0303
10.8888 | -95.0304
10.8892 | -95.0306
10.8895 | -95.0307
10.8898 | -95.0307
10.8902 | -95.0313
10.8905 | -95.0318
10.8909 | -95.0323
10.8912 | -95.0327
10.8915 | -95.0331
10.8919 | -95.0333
10.8922 | -95.0338
10.8926 | -95.034
10.8929 | -95.0343
10.8932 | -95.035
10.8936 | -95.035
10.8939 | -95.0357
10.8943 | -95.0359
10.8946 | -95.0347
10.8949 | -95.0344
10.8953 | -95.035
10.8956 | -95.0354
10.896 | -95.0349
10.8963 | -95.0348
10.8966 | -95.0354
10.897 | -95.0359
10.8973 | -95.0368
10.8977 | -95.0372
10.898 | -95.0378
10.8983 | -95.0385
10.8987 | -95.0388
10.899 | -95.0395
10.8994 | -95.04
10.8997 | -95.0403
10.9 | -95.0409
10.9004 | -95.0414
10.9007 | -95.042
10.9011 | -95.043
10.9014 | -95.044
10.9017 | -95.0451
10.9021 | -95.0456
10.9024 | -95.0465
10.9028 | -95.047
10.9031 | -95.0474
10.9035 | -95.0481
10.9038 | -95.0488
10.9041 | -95.0493
10.9045 | -95.0499
10.9048 | -95.0503
10.9052 | -95.0504
10.9055 | -95.0506
10.9058 | -95.051
10.9062 | -95.0515
10.9065 | -95.052
10.9069 | -95.052
10.9072 | -95.0517
10.9075 | -95.0517
10.9079 | -95.0529
10.9082 | -95.0538
10.9086 | -95.0543
10.9089 | -95.0548
10.9092 | -95.0553
10.9096 | -95.0554
10.9099 | -95.0558
10.9103 | -95.0564
10.9106 | -95.0573
10.9109 | -95.0582
10.9113 | -95.0586
10.9116 | -95.0591
10.912 | -95.0596
10.9123 | -95.0601
10.9126 | -95.0605
10.913 | -95.0611
10.9133 | -95.0615
10.9137 | -95.0621
10.914 | -95.0623
10.9143 | -95.0633
10.9147 | -95.0638
10.915 | -95.0643
10.9154 | -95.0646
10.9157 | -95.065
10.916 | -95.0656
10.9164 | -95.0663
10.9167 | -95.0665
10.9171 | -95.0666
10.9174 | -95.067
10.9178 | -95.0675
10.9181 | -95.0682
10.9184 | -95.0689
10.9188 | -95.0694
10.9191 | -95.07
10.9195 | -95.0702
10.9198 | -95.0705
10.9201 | -95.0713
10.9205 | -95.0719
10.9208 | -95.0723
10.9212 | -95.073
10.9215 | -95.0732
10.9218 | -95.0734
10.9222 | -95.074
10.9225 | -95.0746
10.9229 | -95.0754
10.9232 | -95.0762
10.9235 | -95.0767
10.9239 | -95.0777
10.9242 | -95.0783
10.9246 | -95.0786
10.9249 | -95.0794
10.9252 | -95.08
10.9256 | -95.0808
10.9259 | -95.0808
10.9263 | -95.0811
10.9266 | -95.0815
10.9269 | -95.0825
10.9273 | -95.0823
10.9276 | -95.0827
10.928 | -95.0832
10.9283 | -95.0836
10.9286 | -95.0843
10.929 | -95.0852
10.9293 | -95.0858
10.9297 | -95.0863
10.93 | -95.0872
10.9304 | -95.0878
10.9307 | -95.088
10.931 | -95.0881
10.9314 | -95.0886
10.9317 | -95.0891
10.9321 | -95.0896
10.9324 | -95.0904
10.9327 | -95.091
10.9331 | -95.0915
10.9334 | -95.0916
10.9338 | -95.0927
10.9341 | -95.0935
10.9344 | -95.0942
10.9348 | -95.0949
10.9351 | -95.0955
10.9355 | -95.0962
10.9358 | -95.0968
10.9361 | -95.0974
10.9365 | -95.0976
10.9368 | -95.0978
10.9372 | -95.0984
10.9375 | -95.099
10.9378 | -95.0997
10.9382 | -95.1003
10.9385 | -95.1011
10.9389 | -95.1017
10.9392 | -95.1025
10.9395 | -95.1035
10.9399 | -95.1044
10.9402 | -95.105
10.9406 | -95.1055
10.9409 | -95.1059
10.9412 | -95.1062
10.9416 | -95.1067
10.9419 | -95.1074
10.9423 | -95.108
10.9426 | -95.1086
10.9429 | -95.1089
10.9433 | -95.1094
10.9436 | -95.1105
10.944 | -95.1111
10.9443 | -95.1113
10.9446 | -95.1114
10.945 | -95.1121
10.9453 | -95.113
10.9457 | -95.1137
10.946 | -95.1142
10.9464 | -95.1145
10.9467 | -95.1149
10.947 | -95.115
10.9474 | -95.1154
10.9477 | -95.1162
10.9481 | -95.1175
10.9484 | -95.1183
10.9487 | -95.1187
10.9491 | -95.119
10.9494 | -95.1194
10.9498 | -95.1199
10.9501 | -95.1204
10.9504 | -95.1211
10.9508 | -95.1219
10.9511 | -95.1225
10.9515 | -95.1232
10.9518 | -95.1238
10.9521 | -95.1232
10.9525 | -95.1224
10.9528 | -95.1229
10.9532 | -95.1232
10.9535 | -95.124
10.9538 | -95.1246
10.9542 | -95.1261
10.9545 | -95.1272
10.9549 | -95.1281
10.9552 | -95.1297
10.9555 | -95.131
10.9559 | -95.1309
10.9562 | -95.1309
10.9566 | -95.1317
10.9569 | -95.1325
10.9572 | -95.1331
10.9576 | -95.1336
10.9579 | -95.1342
10.9583 | -95.1348
10.9586 | -95.1355
10.959 | -95.1356
10.9593 | -95.1359
10.9596 | -95.1368
10.96 | -95.1374
10.9603 | -95.1379
10.9607 | -95.1384
10.961 | -95.139
10.9613 | -95.1393
10.9617 | -95.1399
10.962 | -95.1404
10.9624 | -95.1411
10.9627 | -95.1419
10.963 | -95.1427
10.9634 | -95.1434
10.9637 | -95.1438
10.9641 | -95.1445
10.9644 | -95.145
10.9647 | -95.1459
10.9651 | -95.1468
10.9654 | -95.1476
10.9658 | -95.1483
10.9661 | -95.1489
10.9664 | -95.1497
10.9668 | -95.1506
10.9671 | -95.1511
10.9675 | -95.1512
10.9678 | -95.1519
10.9681 | -95.1526
10.9685 | -95.1533
10.9688 | -95.1538
10.9692 | -95.1542
10.9695 | -95.1549
10.9698 | -95.1553
10.9702 | -95.1558
10.9705 | -95.1564
10.9709 | -95.157
10.9712 | -95.1575
10.9715 | -95.1582
10.9719 | -95.1592
10.9722 | -95.16
10.9726 | -95.1609
10.9729 | -95.162
10.9733 | -95.1626
10.9736 | -95.1636
10.9739 | -95.1641
10.9743 | -95.1649
10.9746 | -95.1658
10.975 | -95.1667
10.9753 | -95.1675
10.9756 | -95.1683
10.976 | -95.1688
10.9763 | -95.169
10.9767 | -95.17
10.977 | -95.1709
10.9773 | -95.1714
10.9777 | -95.172
10.978 | -95.1729
10.9784 | -95.1733
10.9787 | -95.1741
10.979 | -95.1748
10.9794 | -95.175
10.9797 | -95.1752
10.9801 | -95.1755
10.9804 | -95.1762
10.9807 | -95.1775
10.9811 | -95.179
10.9814 | -95.1803
10.9818 | -95.181
10.9821 | -95.182
10.9824 | -95.183
10.9828 | -95.1838
10.9831 | -95.1846
10.9835 | -95.1851
10.9838 | -95.1859
10.9841 | -95.1866
10.9845 | -95.1874
10.9848 | -95.1887
10.9852 | -95.1894
10.9855 | -95.19
10.9858 | -95.1906
10.9862 | -95.1915
10.9865 | -95.1924
10.9869 | -95.1931
10.9872 | -95.1937
10.9876 | -95.1942
10.9879 | -95.195
10.9882 | -95.1959
10.9886 | -95.1967
10.9889 | -95.1975
10.9893 | -95.1982
10.9896 | -95.1991
10.9899 | -95.1998
10.9903 | -95.2004
10.9906 | -95.201
10.991 | -95.2016
10.9913 | -95.2021
10.9916 | -95.2028
10.992 | -95.2035
10.9923 | -95.2044
10.9927 | -95.2049
10.993 | -95.2056
10.9933 | -95.2064
10.9937 | -95.2073
10.994 | -95.2081
10.9944 | -95.2089
10.9947 | -95.2094
10.995 | -95.2099
10.9954 | -95.2105
10.9957 | -95.2109
10.9961 | -95.2114
10.9964 | -95.2122
10.9967 | -95.2128
10.9971 | -95.2136
10.9974 | -95.2147
10.9978 | -95.2156
10.9981 | -95.2162
10.9984 | -95.2168
10.9988 | -95.2176
10.9991 | -95.2182
10.9995 | -95.2192
10.9998 | -95.2201
11.0001 | -95.2208
11.0005 | -95.2213
11.0008 | -95.2224
11.0012 | -95.2235
11.0015 | -95.2244
11.0019 | -95.2252
11.0022 | -95.2257
11.0025 | -95.2258
11.0029 | -95.2261
11.0032 | -95.2263
11.0036 | -95.2269
11.0039 | -95.2278
11.0042 | -95.2282
11.0046 | -95.2289
11.0049 | -95.2299
11.0053 | -95.2307
11.0056 | -95.2314
11.0059 | -95.2321
11.0063 | -95.2326
11.0066 | -95.2336
11.007 | -95.2345
11.0073 | -95.2352
11.0076 | -95.2357
11.008 | -95.2365
11.0083 | -95.2368
11.0087 | -95.2375
11.009 | -95.2383
11.0093 | -95.2387
11.0097 | -95.2393
11.01 | -95.2403
11.0104 | -95.2414
11.0107 | -95.2421
11.011 | -95.2428
11.0114 | -95.244
11.0117 | -95.2447
11.0121 | -95.2449
11.0124 | -95.245
11.0127 | -95.2457
11.0131 | -95.2465
11.0134 | -95.247
11.0138 | -95.2473
11.0141 | -95.248
11.0145 | -95.2487
11.0148 | -95.2493
11.0151 | -95.25
11.0155 | -95.2508
11.0158 | -95.2515
11.0162 | -95.2523
11.0165 | -95.2528
11.0168 | -95.2531
11.0172 | -95.2544
11.0175 | -95.2554
11.0179 | -95.2559
11.0182 | -95.2562
11.0185 | -95.2562
11.0189 | -95.2563
11.0192 | -95.257
11.0196 | -95.2579
11.0199 | -95.2586
11.0202 | -95.2592
11.0206 | -95.26
11.0209 | -95.2609
11.0213 | -95.2617
11.0216 | -95.2625
11.0219 | -95.2631
11.0223 | -95.2634
11.0226 | -95.2642
11.023 | -95.2648
11.0233 | -95.2654
11.0236 | -95.266
11.024 | -95.2666
11.0243 | -95.2672
11.0247 | -95.2679
11.025 | -95.2689
11.0253 | -95.2695
11.0257 | -95.2704
11.026 | -95.271
11.0264 | -95.2716
11.0267 | -95.2718
11.027 | -95.2724
11.0274 | -95.2732
11.0277 | -95.2739
11.0281 | -95.2745
11.0284 | -95.2748
11.0288 | -95.2756
11.0291 | -95.276
11.0294 | -95.2769
11.0298 | -95.2777
11.0301 | -95.278
11.0305 | -95.2791
11.0308 | -95.2799
11.0311 | -95.2805
11.0315 | -95.2814
11.0318 | -95.2818
11.0322 | -95.2824
11.0325 | -95.2827
11.0328 | -95.2825
11.0332 | -95.2818
11.0335 | -95.2821
11.0339 | -95.2828
11.0342 | -95.2836
11.0345 | -95.2841
11.0349 | -95.2855
11.0352 | -95.2863
11.0356 | -95.2864
11.0359 | -95.287
11.0362 | -95.2877
11.0366 | -95.2883
11.0369 | -95.2891
11.0373 | -95.2901
11.0376 | -95.2908
11.0379 | -95.2915
11.0383 | -95.2921
11.0386 | -95.2928
11.039 | -95.2935
11.0393 | -95.294
11.0396 | -95.2942
11.04 | -95.2947
11.0403 | -95.2953
11.0407 | -95.2959
11.041 | -95.2963
11.0413 | -95.2967
11.0417 | -95.2975
11.042 | -95.2981
11.0424 | -95.2992
11.0427 | -95.3005
11.0431 | -95.3015
11.0434 | -95.3021
11.0437 | -95.3022
11.0441 | -95.303
11.0444 | -95.3035
11.0448 | -95.304
11.0451 | -95.3044
11.0454 | -95.3048
11.0458 | -95.305
11.0461 | -95.3056
11.0465 | -95.3063
11.0468 | -95.307
11.0471 | -95.3078
11.0475 | -95.3082
11.0478 | -95.3087
11.0482 | -95.3091
11.0485 | -95.3095
11.0488 | -95.3098
11.0492 | -95.31
11.0495 | -95.3106
11.0499 | -95.3112
11.0502 | -95.3114
11.0505 | -95.3111
11.0509 | -95.3117
11.0512 | -95.3127
11.0516 | -95.3127
11.0519 | -95.313
11.0522 | -95.3134
11.0526 | -95.3141
11.0529 | -95.3146
11.0533 | -95.3158
11.0536 | -95.3165
11.0539 | -95.3167
11.0543 | -95.3173
11.0546 | -95.3181
11.055 | -95.3189
11.0553 | -95.3197
11.0556 | -95.3203
11.056 | -95.3207
11.0563 | -95.3218
11.0567 | -95.3229
11.057 | -95.3239
11.0574 | -95.3247
11.0577 | -95.3249
11.058 | -95.3252
11.0584 | -95.3256
11.0587 | -95.3257
11.0591 | -95.3262
11.0594 | -95.3266
11.0597 | -95.3268
11.0601 | -95.3273
11.0604 | -95.3287
11.0608 | -95.3296
11.0611 | -95.3303
11.0614 | -95.3308
11.0618 | -95.332
11.0621 | -95.3332
11.0625 | -95.334
11.0628 | -95.3344
11.0631 | -95.3349
11.0635 | -95.3353
11.0638 | -95.3351
11.0642 | -95.3355
11.0645 | -95.3362
11.0648 | -95.3366
11.0652 | -95.3371
11.0655 | -95.3372
11.0659 | -95.3372
11.0662 | -95.3375
11.0665 | -95.3379
11.0669 | -95.3385
11.0672 | -95.3392
11.0676 | -95.3395
11.0679 | -95.3398
11.0682 | -95.3408
11.0686 | -95.3412
11.0689 | -95.3409
11.0693 | -95.341
11.0696 | -95.3416
11.07 | -95.342
11.0703 | -95.3424
11.0706 | -95.343
11.071 | -95.3445
11.0713 | -95.345
11.0717 | -95.3457
11.072 | -95.3461
11.0723 | -95.3467
11.0727 | -95.3473
11.073 | -95.3479
11.0734 | -95.3481
11.0737 | -95.3486
11.074 | -95.3495
11.0744 | -95.3508
11.0747 | -95.3509
11.0751 | -95.3518
11.0754 | -95.3531
11.0757 | -95.3535
11.0761 | -95.3541
11.0764 | -95.3546
11.0768 | -95.3551
11.0771 | -95.3558
11.0774 | -95.3561
11.0778 | -95.357
11.0781 | -95.3588
11.0785 | -95.36
11.0788 | -95.3606
11.0791 | -95.3612
11.0795 | -95.3616
11.0798 | -95.3619
11.0802 | -95.3624
11.0805 | -95.363
11.0808 | -95.3636
11.0812 | -95.3641
11.0815 | -95.3647
11.0819 | -95.3651
11.0822 | -95.3653
11.0825 | -95.3654
11.0829 | -95.3656
11.0832 | -95.366
11.0836 | -95.3663
11.0839 | -95.3665
11.0843 | -95.3662
11.0846 | -95.3661
11.0849 | -95.3662
11.0853 | -95.3663
11.0856 | -95.3671
11.086 | -95.3675
11.0863 | -95.3672
11.0866 | -95.3675
11.087 | -95.3679
11.0873 | -95.3684
11.0877 | -95.3688
11.088 | -95.3692
11.0883 | -95.3688
11.0887 | -95.3691
11.089 | -95.3695
11.0894 | -95.3695
11.0897 | -95.3691
11.09 | -95.3689
11.0904 | -95.3688
11.0907 | -95.3691
11.0911 | -95.3695
11.0914 | -95.3712
11.0917 | -95.3731
11.0921 | -95.3736
11.0924 | -95.3742
11.0928 | -95.3745
11.0931 | -95.3745
11.0934 | -95.3748
11.0938 | -95.3752
11.0941 | -95.3749
11.0945 | -95.3752
11.0948 | -95.3754
11.0951 | -95.3756
11.0955 | -95.376
11.0958 | -95.3758
11.0962 | -95.3748
11.0965 | -95.3743
11.0968 | -95.3748
11.0972 | -95.3752
11.0975 | -95.3756
11.0979 | -95.3761
11.0982 | -95.3765
11.0986 | -95.3773
11.0989 | -95.3776
11.0992 | -95.3777
11.0996 | -95.3774
11.0999 | -95.3776
11.1003 | -95.378
11.1006 | -95.3785
11.1009 | -95.3781
11.1013 | -95.3779
11.1016 | -95.3781
11.102 | -95.3784
11.1023 | -95.378
11.1026 | -95.378
11.103 | -95.3785
11.1033 | -95.3784
11.1037 | -95.3782
11.104 | -95.3777
11.1043 | -95.3775
11.1047 | -95.3782
11.105 | -95.3786
11.1054 | -95.3793
11.1057 | -95.3803
11.106 | -95.3807
11.1064 | -95.381
11.1067 | -95.3812
11.1071 | -95.3818
11.1074 | -95.3826
11.1077 | -95.3831
11.1081 | -95.3837
11.1084 | -95.3842
11.1088 | -95.3845
11.1091 | -95.3845
11.1094 | -95.3841
11.1098 | -95.3836
11.1101 | -95.3832
11.1105 | -95.3833
11.1108 | -95.3829
11.1111 | -95.3835
11.1115 | -95.3842
11.1118 | -95.3843
11.1122 | -95.3848
11.1125 | -95.3851
11.1129 | -95.3853
11.1132 | -95.3853
11.1135 | -95.3853
11.1139 | -95.386
11.1142 | -95.3871
11.1146 | -95.3875
11.1149 | -95.3878
11.1152 | -95.388
11.1156 | -95.3882
11.1159 | -95.3885
11.1163 | -95.3886
11.1166 | -95.3883
11.1169 | -95.3884
11.1173 | -95.3886
11.1176 | -95.3891
11.118 | -95.3897
11.1183 | -95.3898
11.1186 | -95.3903
11.119 | -95.3915
11.1193 | -95.3923
11.1197 | -95.3929
11.12 | -95.3931
11.1203 | -95.3933
11.1207 | -95.3935
11.121 | -95.3936
11.1214 | -95.3936
11.1217 | -95.3935
11.122 | -95.3933
11.1224 | -95.3935
11.1227 | -95.3935
11.1231 | -95.3935
11.1234 | -95.3936
11.1237 | -95.394
11.1241 | -95.3945
11.1244 | -95.3945
11.1248 | -95.3946
11.1251 | -95.3949
11.1254 | -95.3949
11.1258 | -95.395
11.1261 | -95.3952
11.1265 | -95.3954
11.1268 | -95.3957
11.1272 | -95.3959
11.1275 | -95.3961
11.1278 | -95.3963
11.1282 | -95.3967
11.1285 | -95.397
11.1289 | -95.3972
11.1292 | -95.3976
11.1295 | -95.3976
11.1299 | -95.3979
11.1302 | -95.3979
11.1306 | -95.3983
11.1309 | -95.3984
11.1312 | -95.3985
11.1316 | -95.3987
11.1319 | -95.3986
11.1323 | -95.3981
11.1326 | -95.3977
11.1329 | -95.3979
11.1333 | -95.3979
11.1336 | -95.3981
11.134 | -95.3982
11.1343 | -95.3985
11.1346 | -95.399
11.135 | -95.399
11.1353 | -95.3989
11.1357 | -95.3991
11.136 | -95.3989
11.1363 | -95.3983
11.1367 | -95.3979
11.137 | -95.3975
11.1374 | -95.3974
11.1377 | -95.3972
11.138 | -95.3974
11.1384 | -95.3984
11.1387 | -95.3988
11.1391 | -95.3989
11.1394 | -95.3992
11.1397 | -95.3994
11.1401 | -95.3995
11.1404 | -95.3999
11.1408 | -95.4002
11.1411 | -95.401
11.1415 | -95.4018
11.1418 | -95.4026
11.1421 | -95.4027
11.1425 | -95.4034
11.1428 | -95.4042
11.1432 | -95.4045
11.1435 | -95.4048
11.1438 | -95.4052
11.1442 | -95.4056
11.1445 | -95.4056
11.1449 | -95.4055
11.1452 | -95.4054
11.1455 | -95.4057
11.1459 | -95.406
11.1462 | -95.4059
11.1466 | -95.4062
11.1469 | -95.4064
11.1472 | -95.4067
11.1476 | -95.4071
11.1479 | -95.4074
11.1483 | -95.4074
11.1486 | -95.4077
11.1489 | -95.4087
11.1493 | -95.4094
11.1496 | -95.4093
11.15 | -95.409
11.1503 | -95.4091
11.1506 | -95.4093
11.151 | -95.4095
11.1513 | -95.4097
11.1517 | -95.4095
11.152 | -95.4093
11.1523 | -95.4093
11.1527 | -95.4092
11.153 | -95.4094
11.1534 | -95.4097
11.1537 | -95.4094
11.154 | -95.4094
11.1544 | -95.4094
11.1547 | -95.4096
11.1551 | -95.4095
11.1554 | -95.4092
11.1558 | -95.4095
11.1561 | -95.4097
11.1564 | -95.4099
11.1568 | -95.4097
11.1571 | -95.4099
11.1575 | -95.4105
11.1578 | -95.4108
11.1581 | -95.4107
11.1585 | -95.411
11.1588 | -95.4113
11.1592 | -95.4118
11.1595 | -95.4116
11.1598 | -95.4114
11.1602 | -95.4114
11.1605 | -95.4117
11.1609 | -95.4123
11.1612 | -95.4127
11.1615 | -95.4133
11.1619 | -95.4133
11.1622 | -95.4133
11.1626 | -95.4136
11.1629 | -95.4137
11.1632 | -95.4136
11.1636 | -95.414
11.1639 | -95.4141
11.1643 | -95.4138
11.1646 | -95.4138
11.1649 | -95.4139
11.1653 | -95.4144
11.1656 | -95.4147
11.166 | -95.4154
11.1663 | -95.4159
11.1666 | -95.4161
11.167 | -95.4161
11.1673 | -95.4162
11.1677 | -95.4161
11.168 | -95.416
11.1684 | -95.4159
11.1687 | -95.4161
11.169 | -95.4163
11.1694 | -95.4164
11.1697 | -95.4164
11.1701 | -95.4168
11.1704 | -95.4169
11.1707 | -95.4174
11.1711 | -95.4177
11.1714 | -95.4179
11.1718 | -95.418
11.1721 | -95.418
11.1724 | -95.4179
11.1728 | -95.4179
11.1731 | -95.418
11.1735 | -95.4183
11.1738 | -95.4185
11.1741 | -95.4184
11.1745 | -95.4187
11.1748 | -95.4193
11.1752 | -95.4196
11.1755 | -95.4196
11.1758 | -95.4197
11.1762 | -95.4198
11.1765 | -95.4199
11.1769 | -95.4199
11.1772 | -95.4202
11.1775 | -95.4202
11.1779 | -95.4204
11.1782 | -95.4205
11.1786 | -95.4206
11.1789 | -95.4209
11.1792 | -95.4202
11.1796 | -95.4197
11.1799 | -95.4196
11.1803 | -95.42
11.1806 | -95.4205
11.1809 | -95.4206
11.1813 | -95.4208
11.1816 | -95.421
11.182 | -95.4209
11.1823 | -95.4209
11.1827 | -95.4209
11.183 | -95.4208
11.1833 | -95.4205
11.1837 | -95.42
11.184 | -95.4197
11.1844 | -95.4202
11.1847 | -95.4202
11.185 | -95.42
11.1854 | -95.4201
11.1857 | -95.4204
11.1861 | -95.4209
11.1864 | -95.4209
11.1867 | -95.4209
11.1871 | -95.4211
11.1874 | -95.4216
11.1878 | -95.4216
11.1881 | -95.4215
11.1884 | -95.4203
11.1888 | -95.4194
11.1891 | -95.4195
11.1895 | -95.4194
11.1898 | -95.4194
11.1901 | -95.4192
11.1905 | -95.4191
11.1908 | -95.4191
11.1912 | -95.4189
11.1915 | -95.4188
11.1918 | -95.419
11.1922 | -95.42
11.1925 | -95.4216
11.1929 | -95.4232
11.1932 | -95.4237
11.1935 | -95.4237
11.1939 | -95.424
11.1942 | -95.4247
11.1946 | -95.425
11.1949 | -95.4257
11.1952 | -95.4264
11.1956 | -95.4263
11.1959 | -95.4262
11.1963 | -95.4263
11.1966 | -95.4264
11.197 | -95.4268
11.1973 | -95.4269
11.1976 | -95.427
11.198 | -95.4273
11.1983 | -95.4274
11.1987 | -95.4272
11.199 | -95.4272
11.1993 | -95.4271
11.1997 | -95.427
11.2 | -95.4268
11.2004 | -95.426
11.2007 | -95.4256
11.201 | -95.4254
11.2014 | -95.425
11.2017 | -95.4248
11.2021 | -95.4251
11.2024 | -95.4258
11.2027 | -95.4263
11.2031 | -95.4265
11.2034 | -95.4266
11.2038 | -95.4263
11.2041 | -95.4261
11.2044 | -95.4261
11.2048 | -95.426
11.2051 | -95.4262
11.2055 | -95.4262
11.2058 | -95.426
11.2061 | -95.4261
11.2065 | -95.4259
11.2068 | -95.4256
11.2072 | -95.4256
11.2075 | -95.4256
11.2078 | -95.4256
11.2082 | -95.4248
11.2085 | -95.4241
11.2089 | -95.4237
11.2092 | -95.4237
11.2096 | -95.4237
11.2099 | -95.4232
11.2102 | -95.423
11.2106 | -95.423
11.2109 | -95.423
11.2113 | -95.4229
11.2116 | -95.4225
11.2119 | -95.4224
11.2123 | -95.4223
11.2126 | -95.4223
11.213 | -95.4216
11.2133 | -95.4214
11.2136 | -95.4214
11.214 | -95.4213
11.2143 | -95.4212
11.2147 | -95.4211
11.215 | -95.4212
11.2153 | -95.4207
11.2157 | -95.4202
11.216 | -95.4202
11.2164 | -95.4201
11.2167 | -95.42
11.217 | -95.4202
11.2174 | -95.4201
11.2177 | -95.4203
11.2181 | -95.4204
11.2184 | -95.4204
11.2187 | -95.4197
11.2191 | -95.4196
11.2194 | -95.4198
11.2198 | -95.4196
11.2201 | -95.4195
11.2204 | -95.4194
11.2208 | -95.4191
11.2211 | -95.4186
11.2215 | -95.4183
11.2218 | -95.418
11.2221 | -95.418
11.2225 | -95.418
11.2228 | -95.4181
11.2232 | -95.4182
11.2235 | -95.418
11.2239 | -95.418
11.2242 | -95.4194
11.2245 | -95.4203
11.2249 | -95.4203
11.2252 | -95.4202
11.2256 | -95.4204
11.2259 | -95.4208
11.2262 | -95.4204
11.2266 | -95.42
11.2269 | -95.4198
11.2273 | -95.4202
11.2276 | -95.42
11.2279 | -95.4197
11.2283 | -95.4201
11.2286 | -95.42
11.229 | -95.4203
11.2293 | -95.4201
11.2296 | -95.4198
11.23 | -95.4197
11.2303 | -95.4196
11.2307 | -95.4195
11.231 | -95.4194
11.2313 | -95.4194
11.2317 | -95.4191
11.232 | -95.419
11.2324 | -95.419
11.2327 | -95.4186
11.233 | -95.4185
11.2334 | -95.4183
11.2337 | -95.4183
11.2341 | -95.4184
11.2344 | -95.4184
11.2347 | -95.4184
11.2351 | -95.4185
11.2354 | -95.4186
11.2358 | -95.4188
11.2361 | -95.4191
11.2364 | -95.4192
11.2368 | -95.4184
11.2371 | -95.4171
11.2375 | -95.4162
11.2378 | -95.4164
11.2382 | -95.4163
11.2385 | -95.4166
11.2388 | -95.4164
11.2392 | -95.4164
11.2395 | -95.4168
11.2399 | -95.4167
11.2402 | -95.4164
11.2405 | -95.4163
11.2409 | -95.4161
11.2412 | -95.416
11.2416 | -95.4158
11.2419 | -95.4159
11.2422 | -95.4158
11.2426 | -95.4164
11.2429 | -95.4168
11.2433 | -95.4166
11.2436 | -95.4166
11.2439 | -95.4166
11.2443 | -95.4163
11.2446 | -95.4159
11.245 | -95.4151
11.2453 | -95.4149
11.2456 | -95.4148
11.246 | -95.4142
11.2463 | -95.4142
11.2467 | -95.4146
11.247 | -95.4146
11.2473 | -95.4144
11.2477 | -95.4141
11.248 | -95.414
11.2484 | -95.4139
11.2487 | -95.4136
11.249 | -95.4135
11.2494 | -95.4142
11.2497 | -95.4148
11.2501 | -95.4147
11.2504 | -95.4147
11.2507 | -95.4147
11.2511 | -95.4148
11.2514 | -95.4149
11.2518 | -95.415
11.2521 | -95.4149
11.2525 | -95.4149
11.2528 | -95.4154
11.2531 | -95.4154
11.2535 | -95.4153
11.2538 | -95.4153
11.2542 | -95.4153
11.2545 | -95.4151
11.2548 | -95.4151
11.2552 | -95.415
11.2555 | -95.415
11.2559 | -95.415
11.2562 | -95.4147
11.2565 | -95.4148
11.2569 | -95.4146
11.2572 | -95.4143
11.2576 | -95.4142
11.2579 | -95.414
11.2582 | -95.414
11.2586 | -95.4145
11.2589 | -95.4146
11.2593 | -95.4144
11.2596 | -95.4143
11.2599 | -95.4143
11.2603 | -95.4137
11.2606 | -95.4137
11.261 | -95.4136
11.2613 | -95.4138
11.2616 | -95.4139
11.262 | -95.4138
11.2623 | -95.4133
11.2627 | -95.4129
11.263 | -95.4129
11.2633 | -95.4129
11.2637 | -95.4132
11.264 | -95.4136
11.2644 | -95.4135
11.2647 | -95.4132
11.2651 | -95.4134
11.2654 | -95.4134
11.2657 | -95.4135
11.2661 | -95.4138
11.2664 | -95.4138
11.2668 | -95.414
11.2671 | -95.414
11.2674 | -95.4138
11.2678 | -95.4137
11.2681 | -95.4131
11.2685 | -95.4135
11.2688 | -95.4143
11.2691 | -95.4149
11.2695 | -95.4143
11.2698 | -95.4142
11.2702 | -95.4138
11.2705 | -95.4132
11.2708 | -95.4131
11.2712 | -95.413
11.2715 | -95.4129
11.2719 | -95.4128
11.2722 | -95.4123
11.2725 | -95.4125
11.2729 | -95.4126
11.2732 | -95.4123
11.2736 | -95.4121
11.2739 | -95.4121
11.2742 | -95.4125
11.2746 | -95.4119
11.2749 | -95.411
11.2753 | -95.4109
11.2756 | -95.4105
11.2759 | -95.4103
11.2763 | -95.4104
11.2766 | -95.4102
11.277 | -95.4106
11.2773 | -95.411
11.2776 | -95.4112
11.278 | -95.4105
11.2783 | -95.4097
11.2787 | -95.4096
11.279 | -95.4099
11.2794 | -95.4104
11.2797 | -95.4111
11.28 | -95.4115
11.2804 | -95.4122
11.2807 | -95.4138
11.2811 | -95.4151
11.2814 | -95.4156
11.2817 | -95.4161
11.2821 | -95.416
11.2824 | -95.4159
11.2828 | -95.4157
11.2831 | -95.4157
11.2834 | -95.416
11.2838 | -95.4161
11.2841 | -95.416
11.2845 | -95.416
11.2848 | -95.4158
11.2851 | -95.4161
11.2855 | -95.4162
11.2858 | -95.416
11.2862 | -95.4153
11.2865 | -95.4151
11.2868 | -95.4152
11.2872 | -95.415
11.2875 | -95.4152
11.2879 | -95.4152
11.2882 | -95.4154
11.2885 | -95.4153
11.2889 | -95.4152
11.2892 | -95.4153
11.2896 | -95.4156
11.2899 | -95.4155
11.2902 | -95.4154
11.2906 | -95.4153
11.2909 | -95.415
11.2913 | -95.4144
11.2916 | -95.4141
11.2919 | -95.4141
11.2923 | -95.4146
11.2926 | -95.4149
11.293 | -95.4145
11.2933 | -95.414
11.2937 | -95.4138
11.294 | -95.4136
11.2943 | -95.4137
11.2947 | -95.4135
11.295 | -95.4133
11.2954 | -95.4132
11.2957 | -95.4126
11.296 | -95.4121
11.2964 | -95.4123
11.2967 | -95.4122
11.2971 | -95.4119
11.2974 | -95.4122
11.2977 | -95.4121
11.2981 | -95.4126
11.2984 | -95.413
11.2988 | -95.4131
11.2991 | -95.4133
11.2994 | -95.4132
11.2998 | -95.4134
11.3001 | -95.4133
11.3005 | -95.4135
11.3008 | -95.4137
11.3011 | -95.414
11.3015 | -95.4144
11.3018 | -95.4147
11.3022 | -95.4143
11.3025 | -95.4136
11.3028 | -95.4135
11.3032 | -95.4128
11.3035 | -95.4121
11.3039 | -95.412
11.3042 | -95.4123
11.3045 | -95.4126
11.3049 | -95.413
11.3052 | -95.4137
11.3056 | -95.4135
11.3059 | -95.4134
11.3062 | -95.4133
11.3066 | -95.4135
11.3069 | -95.4133
11.3073 | -95.4134
11.3076 | -95.4133
11.308 | -95.4131
11.3083 | -95.413
11.3086 | -95.4128
11.309 | -95.4126
11.3093 | -95.4125
11.3097 | -95.4127
11.31 | -95.4128
11.3103 | -95.4127
11.3107 | -95.4124
11.311 | -95.4122
11.3114 | -95.4122
11.3117 | -95.4121
11.312 | -95.4118
11.3124 | -95.4115
11.3127 | -95.4114
11.3131 | -95.4113
11.3134 | -95.4115
11.3137 | -95.4115
11.3141 | -95.4114
11.3144 | -95.4113
11.3148 | -95.4111
11.3151 | -95.4112
11.3154 | -95.4114
11.3158 | -95.4114
11.3161 | -95.4114
11.3165 | -95.4115
11.3168 | -95.4119
11.3171 | -95.4123
11.3175 | -95.4122
11.3178 | -95.412
11.3182 | -95.412
11.3185 | -95.4119
11.3188 | -95.4122
11.3192 | -95.4122
11.3195 | -95.4123
11.3199 | -95.4118
11.3202 | -95.4114
11.3205 | -95.4113
11.3209 | -95.4111
11.3212 | -95.4108
11.3216 | -95.411
11.3219 | -95.4115
11.3223 | -95.4118
11.3226 | -95.4125
11.3229 | -95.4132
11.3233 | -95.4133
11.3236 | -95.4131
11.324 | -95.4131
11.3243 | -95.4131
11.3246 | -95.4132
11.325 | -95.4137
11.3253 | -95.4136
11.3257 | -95.4137
11.326 | -95.4139
11.3263 | -95.4137
11.3267 | -95.413
11.327 | -95.4128
11.3274 | -95.4127
11.3277 | -95.4128
11.328 | -95.4128
11.3284 | -95.4122
11.3287 | -95.4125
11.3291 | -95.4128
11.3294 | -95.4133
11.3297 | -95.4135
11.3301 | -95.4132
11.3304 | -95.4128
11.3308 | -95.4128
11.3311 | -95.413
11.3314 | -95.4131
11.3318 | -95.4131
11.3321 | -95.4135
11.3325 | -95.4135
11.3328 | -95.4136
11.3331 | -95.4139
11.3335 | -95.4139
11.3338 | -95.4139
11.3342 | -95.4142
11.3345 | -95.4137
11.3348 | -95.4139
11.3352 | -95.4148
11.3355 | -95.415
11.3359 | -95.4152
11.3362 | -95.4154
11.3366 | -95.416
11.3369 | -95.4157
11.3372 | -95.4156
11.3376 | -95.4156
11.3379 | -95.4158
11.3383 | -95.4164
11.3386 | -95.4171
11.3389 | -95.4173
11.3393 | -95.4174
11.3396 | -95.4178
11.34 | -95.4183
11.3403 | -95.4186
11.3406 | -95.4186
11.341 | -95.4181
11.3413 | -95.4176
11.3417 | -95.4177
11.342 | -95.4177
11.3423 | -95.4177
11.3427 | -95.4176
11.343 | -95.4174
11.3434 | -95.417
11.3437 | -95.4168
11.344 | -95.417
11.3444 | -95.4171
11.3447 | -95.4172
11.3451 | -95.4178
11.3454 | -95.4184
11.3457 | -95.4184
11.3461 | -95.4184
11.3464 | -95.4181
11.3468 | -95.4175
11.3471 | -95.4174
11.3474 | -95.4175
11.3478 | -95.4168
11.3481 | -95.4164
11.3485 | -95.4167
11.3488 | -95.4166
11.3491 | -95.4166
11.3495 | -95.4167
11.3498 | -95.4166
11.3502 | -95.4164
11.3505 | -95.416
11.3509 | -95.4158
11.3512 | -95.4159
11.3515 | -95.416
11.3519 | -95.4161
11.3522 | -95.4161
11.3526 | -95.4159
11.3529 | -95.4158
11.3532 | -95.416
11.3536 | -95.4161
11.3539 | -95.4162
11.3543 | -95.4162
11.3546 | -95.4163
11.3549 | -95.4163
11.3553 | -95.4164
11.3556 | -95.4167
11.356 | -95.4173
11.3563 | -95.4173
11.3566 | -95.4173
11.357 | -95.4173
11.3573 | -95.4175
11.3577 | -95.4179
11.358 | -95.4181
11.3583 | -95.4178
11.3587 | -95.4178
11.359 | -95.4177
11.3594 | -95.4177
11.3597 | -95.4176
11.36 | -95.4175
11.3604 | -95.4176
11.3607 | -95.418
11.3611 | -95.4185
11.3614 | -95.4187
11.3617 | -95.4188
11.3621 | -95.4189
11.3624 | -95.419
11.3628 | -95.4191
11.3631 | -95.419
11.3635 | -95.4193
11.3638 | -95.4195
11.3641 | -95.4194
11.3645 | -95.4189
11.3648 | -95.4187
11.3652 | -95.4187
11.3655 | -95.4188
11.3658 | -95.4185
11.3662 | -95.4185
11.3665 | -95.4184
11.3669 | -95.4185
11.3672 | -95.4182
11.3675 | -95.4189
11.3679 | -95.4195
11.3682 | -95.4196
11.3686 | -95.4197
11.3689 | -95.4201
11.3692 | -95.4205
11.3696 | -95.4205
11.3699 | -95.4206
11.3703 | -95.4206
11.3706 | -95.4204
11.3709 | -95.4203
11.3713 | -95.4205
11.3716 | -95.4206
11.372 | -95.4214
11.3723 | -95.4227
11.3726 | -95.423
11.373 | -95.423
11.3733 | -95.423
11.3737 | -95.4231
11.374 | -95.4233
11.3743 | -95.4233
11.3747 | -95.4234
11.375 | -95.4236
11.3754 | -95.4239
11.3757 | -95.4241
11.376 | -95.4244
11.3764 | -95.4255
11.3767 | -95.4261
11.3771 | -95.4261
11.3774 | -95.4264
11.3778 | -95.4265
11.3781 | -95.4267
11.3784 | -95.4268
11.3788 | -95.4268
11.3791 | -95.4271
11.3795 | -95.4273
11.3798 | -95.4275
11.3801 | -95.4275
11.3805 | -95.4276
11.3808 | -95.4275
11.3812 | -95.4276
11.3815 | -95.4276
11.3818 | -95.4277
11.3822 | -95.4279
11.3825 | -95.428
11.3829 | -95.428
11.3832 | -95.4283
11.3835 | -95.4288
11.3839 | -95.4284
11.3842 | -95.4284
11.3846 | -95.4289
11.3849 | -95.4293
11.3852 | -95.4294
11.3856 | -95.4296
11.3859 | -95.4299
11.3863 | -95.4299
11.3866 | -95.4298
11.3869 | -95.4297
11.3873 | -95.4306
11.3876 | -95.4316
11.388 | -95.4321
11.3883 | -95.4326
11.3886 | -95.4329
11.389 | -95.4328
11.3893 | -95.433
11.3897 | -95.4333
11.39 | -95.4335
11.3903 | -95.434
11.3907 | -95.4342
11.391 | -95.4346
11.3914 | -95.4357
11.3917 | -95.4364
11.3921 | -95.4363
11.3924 | -95.4366
11.3927 | -95.4373
11.3931 | -95.4375
11.3934 | -95.4377
11.3938 | -95.4379
11.3941 | -95.4381
11.3944 | -95.4386
11.3948 | -95.4396
11.3951 | -95.4398
11.3955 | -95.4404
11.3958 | -95.4405
11.3961 | -95.4406
11.3965 | -95.4407
11.3968 | -95.4409
11.3972 | -95.4411
11.3975 | -95.4411
11.3978 | -95.4411
11.3982 | -95.4413
11.3985 | -95.4414
11.3989 | -95.4414
11.3992 | -95.4414
11.3995 | -95.4414
11.3999 | -95.4415
11.4002 | -95.4415
11.4006 | -95.4417
11.4009 | -95.442
11.4012 | -95.4425
11.4016 | -95.4426
11.4019 | -95.4426
11.4023 | -95.4429
11.4026 | -95.4433
11.4029 | -95.4435
11.4033 | -95.4434
11.4036 | -95.4437
11.404 | -95.4438
11.4043 | -95.4438
11.4047 | -95.4441
11.405 | -95.4443
11.4053 | -95.4445
11.4057 | -95.4446
11.406 | -95.4448
11.4064 | -95.4453
11.4067 | -95.4453
11.407 | -95.4459
11.4074 | -95.4465
11.4077 | -95.4467
11.4081 | -95.4467
11.4084 | -95.4471
11.4087 | -95.4475
11.4091 | -95.4473
11.4094 | -95.4472
11.4098 | -95.4473
11.4101 | -95.4474
11.4104 | -95.4482
11.4108 | -95.4487
11.4111 | -95.4488
11.4115 | -95.4482
11.4118 | -95.4481
11.4121 | -95.4482
11.4125 | -95.4481
11.4128 | -95.4485
11.4132 | -95.4486
11.4135 | -95.4488
11.4138 | -95.4493
11.4142 | -95.4494
11.4145 | -95.4495
11.4149 | -95.4494
11.4152 | -95.4485
11.4155 | -95.4486
11.4159 | -95.4487
11.4162 | -95.4486
11.4166 | -95.4488
11.4169 | -95.4496
11.4172 | -95.45
11.4176 | -95.4502
11.4179 | -95.4499
11.4183 | -95.4497
11.4186 | -95.4498
11.419 | -95.45
11.4193 | -95.4501
11.4196 | -95.4501
11.42 | -95.4501
11.4203 | -95.4504
11.4207 | -95.4507
11.421 | -95.4507
11.4213 | -95.4511
11.4217 | -95.4512
11.422 | -95.4512
11.4224 | -95.4513
11.4227 | -95.4511
11.423 | -95.4509
11.4234 | -95.4513
11.4237 | -95.4515
11.4241 | -95.4519
11.4244 | -95.4524
11.4247 | -95.4531
11.4251 | -95.4536
11.4254 | -95.4538
11.4258 | -95.4542
11.4261 | -95.4545
11.4264 | -95.4547
11.4268 | -95.4548
11.4271 | -95.455
11.4275 | -95.4553
11.4278 | -95.4555
11.4281 | -95.4558
11.4285 | -95.4561
11.4288 | -95.4564
11.4292 | -95.4563
11.4295 | -95.4565
11.4298 | -95.4565
11.4302 | -95.4571
11.4305 | -95.4573
11.4309 | -95.4575
11.4312 | -95.4576
11.4315 | -95.4575
11.4319 | -95.4576
11.4322 | -95.4579
11.4326 | -95.458
11.4329 | -95.4582
11.4333 | -95.4585
11.4336 | -95.4589
11.4339 | -95.4592
11.4343 | -95.4594
11.4346 | -95.4597
11.435 | -95.4597
11.4353 | -95.4598
11.4356 | -95.4606
11.436 | -95.4611
11.4363 | -95.4614
11.4367 | -95.4617
11.437 | -95.4619
11.4373 | -95.4619
11.4377 | -95.4624
11.438 | -95.4626
11.4384 | -95.4623
11.4387 | -95.4621
11.439 | -95.4619
11.4394 | -95.462
11.4397 | -95.4625
11.4401 | -95.4624
11.4404 | -95.4625
11.4407 | -95.4626
11.4411 | -95.4629
11.4414 | -95.4632
11.4418 | -95.4635
11.4421 | -95.4638
11.4424 | -95.464
11.4428 | -95.4643
11.4431 | -95.4647
11.4435 | -95.4651
11.4438 | -95.4654
11.4441 | -95.4657
11.4445 | -95.4661
11.4448 | -95.467
11.4452 | -95.4674
11.4455 | -95.4676
11.4458 | -95.4676
11.4462 | -95.4678
11.4465 | -95.4682
11.4469 | -95.4688
11.4472 | -95.469
11.4476 | -95.4689
11.4479 | -95.4693
11.4482 | -95.4709
11.4486 | -95.4723
11.4489 | -95.4727
11.4493 | -95.4728
11.4496 | -95.473
11.4499 | -95.4733
11.4503 | -95.4736
11.4506 | -95.4739
11.451 | -95.4742
11.4513 | -95.4747
11.4516 | -95.4753
11.452 | -95.4758
11.4523 | -95.476
11.4527 | -95.4762
11.453 | -95.4765
11.4533 | -95.4769
11.4537 | -95.4776
11.454 | -95.4777
11.4544 | -95.478
11.4547 | -95.4782
11.455 | -95.4784
11.4554 | -95.4786
11.4557 | -95.4793
11.4561 | -95.48
11.4564 | -95.4799
11.4567 | -95.4801
11.4571 | -95.4805
11.4574 | -95.4809
11.4578 | -95.4813
11.4581 | -95.4816
11.4584 | -95.4817
11.4588 | -95.4818
11.4591 | -95.4821
11.4595 | -95.4828
11.4598 | -95.4834
11.4601 | -95.4838
11.4605 | -95.4837
11.4608 | -95.4833
11.4612 | -95.4834
11.4615 | -95.4835
11.4619 | -95.4839
11.4622 | -95.4843
11.4625 | -95.4852
11.4629 | -95.4864
11.4632 | -95.487
11.4636 | -95.4873
11.4639 | -95.4876
11.4642 | -95.4883
11.4646 | -95.4885
11.4649 | -95.4891
11.4653 | -95.4897
11.4656 | -95.49
11.4659 | -95.4905
11.4663 | -95.4909
11.4666 | -95.4915
11.467 | -95.4923
11.4673 | -95.4923
11.4676 | -95.4922
11.468 | -95.4926
11.4683 | -95.493
11.4687 | -95.4934
11.469 | -95.4936
11.4693 | -95.4939
11.4697 | -95.494
11.47 | -95.4944
11.4704 | -95.495
11.4707 | -95.4953
11.471 | -95.4954
11.4714 | -95.4955
11.4717 | -95.4959
11.4721 | -95.496
11.4724 | -95.4965
11.4727 | -95.4969
11.4731 | -95.4973
11.4734 | -95.4977
11.4738 | -95.498
11.4741 | -95.4977
11.4744 | -95.4974
11.4748 | -95.4978
11.4751 | -95.4984
11.4755 | -95.4988
11.4758 | -95.4991
11.4762 | -95.4995
11.4765 | -95.4999
11.4768 | -95.4998
11.4772 | -95.4993
11.4775 | -95.4992
11.4779 | -95.4994
11.4782 | -95.4995
11.4785 | -95.4994
11.4789 | -95.4996
11.4792 | -95.4997
11.4796 | -95.5003
11.4799 | -95.5011
11.4802 | -95.5013
11.4806 | -95.5013
11.4809 | -95.5015
11.4813 | -95.5023
11.4816 | -95.5032
11.4819 | -95.5035
11.4823 | -95.5036
11.4826 | -95.504
11.483 | -95.5045
11.4833 | -95.5049
11.4836 | -95.5051
11.484 | -95.5048
11.4843 | -95.5048
11.4847 | -95.5052
11.485 | -95.5058
11.4853 | -95.5064
11.4857 | -95.5068
11.486 | -95.5077
11.4864 | -95.5084
11.4867 | -95.5084
11.487 | -95.5086
11.4874 | -95.509
11.4877 | -95.5093
11.4881 | -95.5097
11.4884 | -95.5104
11.4887 | -95.5111
11.4891 | -95.5116
11.4894 | -95.5118
11.4898 | -95.5124
11.4901 | -95.5131
11.4905 | -95.5134
11.4908 | -95.5138
11.4911 | -95.5142
11.4915 | -95.5145
11.4918 | -95.515
11.4922 | -95.5159
11.4925 | -95.5166
11.4928 | -95.517
11.4932 | -95.5175
11.4935 | -95.5181
11.4939 | -95.5188
11.4942 | -95.5196
11.4945 | -95.5199
11.4949 | -95.5204
11.4952 | -95.5209
11.4956 | -95.5208
11.4959 | -95.5212
11.4962 | -95.5214
11.4966 | -95.5214
11.4969 | -95.5219
11.4973 | -95.5228
11.4976 | -95.5233
11.4979 | -95.5237
11.4983 | -95.524
11.4986 | -95.5252
11.499 | -95.5258
11.4993 | -95.5262
11.4996 | -95.5265
11.5 | -95.5266
11.5003 | -95.5264
11.5007 | -95.5268
11.501 | -95.5273
11.5013 | -95.5276
11.5017 | -95.5278
11.502 | -95.5282
11.5024 | -95.5283
11.5027 | -95.5282
11.5031 | -95.5283
11.5034 | -95.5283
11.5037 | -95.5289
11.5041 | -95.5293
11.5044 | -95.5297
11.5048 | -95.5301
11.5051 | -95.5307
11.5054 | -95.5316
11.5058 | -95.5321
11.5061 | -95.5321
11.5065 | -95.5325
11.5068 | -95.5335
11.5071 | -95.5346
11.5075 | -95.5352
11.5078 | -95.5357
11.5082 | -95.5361
11.5085 | -95.5364
11.5088 | -95.5365
11.5092 | -95.5366
11.5095 | -95.5369
11.5099 | -95.5372
11.5102 | -95.5377
11.5105 | -95.5379
11.5109 | -95.5383
11.5112 | -95.5375
11.5116 | -95.5373
11.5119 | -95.5375
11.5122 | -95.538
11.5126 | -95.5383
11.5129 | -95.5385
11.5133 | -95.5385
11.5136 | -95.5388
11.5139 | -95.5394
11.5143 | -95.5398
11.5146 | -95.5402
11.515 | -95.541
11.5153 | -95.5418
11.5156 | -95.5428
11.516 | -95.544
11.5163 | -95.5443
11.5167 | -95.5445
11.517 | -95.5452
11.5174 | -95.5454
11.5177 | -95.5458
11.518 | -95.5462
11.5184 | -95.5459
11.5187 | -95.5451
11.5191 | -95.5453
11.5194 | -95.5461
11.5197 | -95.5468
11.5201 | -95.5469
11.5204 | -95.5475
11.5208 | -95.5478
11.5211 | -95.548
11.5214 | -95.5477
11.5218 | -95.5475
11.5221 | -95.5472
11.5225 | -95.5474
11.5228 | -95.5477
11.5231 | -95.5485
11.5235 | -95.5489
11.5238 | -95.5493
11.5242 | -95.5496
11.5245 | -95.55
11.5248 | -95.5503
11.5252 | -95.5505
11.5255 | -95.551
11.5259 | -95.5514
11.5262 | -95.5518
11.5265 | -95.5525
11.5269 | -95.5533
11.5272 | -95.5539
11.5276 | -95.5542
11.5279 | -95.5541
11.5282 | -95.5543
11.5286 | -95.5547
11.5289 | -95.555
11.5293 | -95.5553
11.5296 | -95.5558
11.5299 | -95.5564
11.5303 | -95.557
11.5306 | -95.5578
11.531 | -95.5585
11.5313 | -95.5591
11.5317 | -95.5594
11.532 | -95.5598
11.5323 | -95.5602
11.5327 | -95.5607
11.533 | -95.5609
11.5334 | -95.5609
11.5337 | -95.5611
11.534 | -95.5616
11.5344 | -95.562
11.5347 | -95.5626
11.5351 | -95.563
11.5354 | -95.5634
11.5357 | -95.5637
11.5361 | -95.5637
11.5364 | -95.5636
11.5368 | -95.5638
11.5371 | -95.5645
11.5374 | -95.5651
11.5378 | -95.5655
11.5381 | -95.566
11.5385 | -95.5664
11.5388 | -95.5665
11.5391 | -95.567
11.5395 | -95.5669
11.5398 | -95.5668
11.5402 | -95.5665
11.5405 | -95.5667
11.5408 | -95.567
11.5412 | -95.5673
11.5415 | -95.5677
11.5419 | -95.5685
11.5422 | -95.5693
11.5425 | -95.5697
11.5429 | -95.5704
11.5432 | -95.5708
11.5436 | -95.5712
11.5439 | -95.5716
11.5442 | -95.5722
11.5446 | -95.5722
11.5449 | -95.5716
11.5453 | -95.5712
11.5456 | -95.5713
11.546 | -95.5715
11.5463 | -95.5722
11.5466 | -95.5727
11.547 | -95.5731
11.5473 | -95.5737
11.5477 | -95.5741
11.548 | -95.5745
11.5483 | -95.5746
11.5487 | -95.5752
11.549 | -95.5758
11.5494 | -95.5768
11.5497 | -95.5776
11.55 | -95.5779
11.5504 | -95.5782
11.5507 | -95.5787
11.5511 | -95.5793
11.5514 | -95.5793
11.5517 | -95.58
11.5521 | -95.581
11.5524 | -95.5815
11.5528 | -95.582
11.5531 | -95.5825
11.5534 | -95.583
11.5538 | -95.5833
11.5541 | -95.5841
11.5545 | -95.5842
11.5548 | -95.5842
11.5551 | -95.5851
11.5555 | -95.5855
11.5558 | -95.5858
11.5562 | -95.586
11.5565 | -95.5856
11.5568 | -95.5859
11.5572 | -95.5862
11.5575 | -95.5868
11.5579 | -95.5874
11.5582 | -95.5875
11.5586 | -95.5874
11.5589 | -95.5882
11.5592 | -95.5885
11.5596 | -95.5888
11.5599 | -95.5891
11.5603 | -95.5895
11.5606 | -95.5898
11.5609 | -95.5903
11.5613 | -95.5903
11.5616 | -95.5906
11.562 | -95.591
11.5623 | -95.5915
11.5626 | -95.592
11.563 | -95.5926
11.5633 | -95.5928
11.5637 | -95.5933
11.564 | -95.5935
11.5643 | -95.594
11.5647 | -95.5944
11.565 | -95.595
11.5654 | -95.5953
11.5657 | -95.5951
11.566 | -95.5955
11.5664 | -95.5956
11.5667 | -95.5955
11.5671 | -95.5956
11.5674 | -95.5957
11.5677 | -95.5958
11.5681 | -95.5964
11.5684 | -95.5969
11.5688 | -95.5971
11.5691 | -95.5967
11.5694 | -95.5965
11.5698 | -95.5968
11.5701 | -95.5968
11.5705 | -95.5963
11.5708 | -95.5963
11.5711 | -95.5967
11.5715 | -95.5971
11.5718 | -95.5963
11.5722 | -95.5963
11.5725 | -95.5967
11.5729 | -95.5971
11.5732 | -95.5975
11.5735 | -95.5982
11.5739 | -95.599
11.5742 | -95.5992
11.5746 | -95.5992
11.5749 | -95.5994
11.5752 | -95.5997
11.5756 | -95.6
11.5759 | -95.6004
11.5763 | -95.6006
11.5766 | -95.6007
11.5769 | -95.6013
11.5773 | -95.6013
11.5776 | -95.6018
11.578 | -95.6021
11.5783 | -95.6027
11.5786 | -95.6031
11.579 | -95.6032
11.5793 | -95.6032
11.5797 | -95.6034
11.58 | -95.6038
11.5803 | -95.6042
11.5807 | -95.6044
11.581 | -95.6044
11.5814 | -95.6052
11.5817 | -95.6058
11.582 | -95.6059
11.5824 | -95.6062
11.5827 | -95.6064
11.5831 | -95.6068
11.5834 | -95.6071
11.5837 | -95.6075
11.5841 | -95.608
11.5844 | -95.6083
11.5848 | -95.6087
11.5851 | -95.6093
11.5854 | -95.6097
11.5858 | -95.61
11.5861 | -95.6104
11.5865 | -95.6114
11.5868 | -95.612
11.5872 | -95.6124
11.5875 | -95.6127
11.5878 | -95.6132
11.5882 | -95.6139
11.5885 | -95.6143
11.5889 | -95.615
11.5892 | -95.6156
11.5895 | -95.6159
11.5899 | -95.6162
11.5902 | -95.6165
11.5906 | -95.6167
11.5909 | -95.6171
11.5912 | -95.6177
11.5916 | -95.6181
11.5919 | -95.6185
11.5923 | -95.619
11.5926 | -95.6193
11.5929 | -95.62
11.5933 | -95.6209
11.5936 | -95.6216
11.594 | -95.6219
11.5943 | -95.6222
11.5946 | -95.6226
11.595 | -95.6231
11.5953 | -95.6233
11.5957 | -95.6237
11.596 | -95.6242
11.5963 | -95.6245
11.5967 | -95.6249
11.597 | -95.6248
11.5974 | -95.6254
11.5977 | -95.6259
11.598 | -95.6262
11.5984 | -95.6267
11.5987 | -95.6268
11.5991 | -95.6269
11.5994 | -95.6271
11.5997 | -95.6275
11.6001 | -95.6281
11.6004 | -95.6283
11.6008 | -95.6285
11.6011 | -95.629
11.6015 | -95.6293
11.6018 | -95.6298
11.6021 | -95.6307
11.6025 | -95.6314
11.6028 | -95.6315
11.6032 | -95.6317
11.6035 | -95.6317
11.6038 | -95.6317
11.6042 | -95.6319
11.6045 | -95.6322
11.6049 | -95.6326
11.6052 | -95.633
11.6055 | -95.6333
11.6059 | -95.6328
11.6062 | -95.6324
11.6066 | -95.6329
11.6069 | -95.6341
11.6072 | -95.6346
11.6076 | -95.635
11.6079 | -95.6351
11.6083 | -95.6351
11.6086 | -95.6352
11.6089 | -95.6357
11.6093 | -95.6359
11.6096 | -95.6357
11.61 | -95.6354
11.6103 | -95.6353
11.6106 | -95.6357
11.611 | -95.6361
11.6113 | -95.6364
11.6117 | -95.6368
11.612 | -95.6374
11.6123 | -95.6379
11.6127 | -95.6376
11.613 | -95.6378
11.6134 | -95.6383
11.6137 | -95.6387
11.6141 | -95.6393
11.6144 | -95.6398
11.6147 | -95.6403
11.6151 | -95.6403
11.6154 | -95.6394
11.6158 | -95.6395
11.6161 | -95.6397
11.6164 | -95.6396
11.6168 | -95.6397
11.6171 | -95.6398
11.6175 | -95.6402
11.6178 | -95.6404
11.6181 | -95.6408
11.6185 | -95.6414
11.6188 | -95.6418
11.6192 | -95.642
11.6195 | -95.6423
11.6198 | -95.6427
11.6202 | -95.6432
11.6205 | -95.6435
11.6209 | -95.6439
11.6212 | -95.6444
11.6215 | -95.6427
11.6219 | -95.6423
11.6222 | -95.6426
11.6226 | -95.6426
11.6229 | -95.6426
11.6232 | -95.643
11.6236 | -95.643
11.6239 | -95.6428
11.6243 | -95.6428
11.6246 | -95.6429
11.6249 | -95.6432
11.6253 | -95.6435
11.6256 | -95.6439
11.626 | -95.6446
11.6263 | -95.6459
11.6266 | -95.6464
11.627 | -95.6467
11.6273 | -95.647
11.6277 | -95.6471
11.628 | -95.6475
11.6283 | -95.6477
11.6287 | -95.6478
11.629 | -95.6476
11.6294 | -95.6478
11.6297 | -95.6483
11.6301 | -95.6486
11.6304 | -95.649
11.6307 | -95.6487
11.6311 | -95.6486
11.6314 | -95.6484
11.6318 | -95.6486
11.6321 | -95.649
11.6324 | -95.6493
11.6328 | -95.6493
11.6331 | -95.6493
11.6335 | -95.6492
11.6338 | -95.6494
11.6341 | -95.65
11.6345 | -95.6506
11.6348 | -95.6509
11.6352 | -95.6512
11.6355 | -95.6518
11.6358 | -95.6524
11.6362 | -95.6527
11.6365 | -95.6536
11.6369 | -95.6547
11.6372 | -95.6551
11.6375 | -95.6547
11.6379 | -95.6548
11.6382 | -95.6551
11.6386 | -95.6552
11.6389 | -95.6551
11.6392 | -95.6551
11.6396 | -95.6554
11.6399 | -95.6559
11.6403 | -95.6563
11.6406 | -95.6566
11.6409 | -95.657
11.6413 | -95.6573
11.6416 | -95.6575
11.642 | -95.6579
11.6423 | -95.6584
11.6427 | -95.6586
11.643 | -95.6589
11.6433 | -95.6593
11.6437 | -95.6596
11.644 | -95.6598
11.6444 | -95.6605
11.6447 | -95.6613
11.645 | -95.6615
11.6454 | -95.6617
11.6457 | -95.6622
11.6461 | -95.6627
11.6464 | -95.6629
11.6467 | -95.6629
11.6471 | -95.663
11.6474 | -95.6629
11.6478 | -95.6627
11.6481 | -95.6625
11.6484 | -95.6622
11.6488 | -95.6619
11.6491 | -95.6618
11.6495 | -95.6618
11.6498 | -95.6623
11.6501 | -95.663
11.6505 | -95.6639
11.6508 | -95.664
11.6512 | -95.6634
11.6515 | -95.6631
11.6518 | -95.6635
11.6522 | -95.6636
11.6525 | -95.6636
11.6529 | -95.6638
11.6532 | -95.664
11.6535 | -95.6642
11.6539 | -95.6645
11.6542 | -95.6651
11.6546 | -95.6652
11.6549 | -95.6651
11.6552 | -95.6649
11.6556 | -95.6639
11.6559 | -95.6635
11.6563 | -95.6644
11.6566 | -95.6644
11.657 | -95.6634
11.6573 | -95.6626
11.6576 | -95.6622
11.658 | -95.662
11.6583 | -95.6622
11.6587 | -95.6624
11.659 | -95.6625
11.6593 | -95.6631
11.6597 | -95.6633
11.66 | -95.6634
11.6604 | -95.6635
11.6607 | -95.6634
11.661 | -95.6636
11.6614 | -95.664
11.6617 | -95.6641
11.6621 | -95.6642
11.6624 | -95.6644
11.6627 | -95.6633
11.6631 | -95.6614
11.6634 | -95.6602
11.6638 | -95.6606
11.6641 | -95.6617
11.6644 | -95.6632
11.6648 | -95.6635
11.6651 | -95.6635
11.6655 | -95.6635
11.6658 | -95.6638
11.6661 | -95.6645
11.6665 | -95.6652
11.6668 | -95.6657
11.6672 | -95.6659
11.6675 | -95.6665
11.6678 | -95.6673
11.6682 | -95.6678
11.6685 | -95.6677
11.6689 | -95.6678
11.6692 | -95.6679
11.6695 | -95.668
11.6699 | -95.6685
11.6702 | -95.669
11.6706 | -95.6689
11.6709 | -95.6693
11.6713 | -95.67
11.6716 | -95.6707
11.6719 | -95.6709
11.6723 | -95.6714
11.6726 | -95.6717
11.673 | -95.6714
11.6733 | -95.6714
11.6736 | -95.6715
11.674 | -95.6717
11.6743 | -95.6717
11.6747 | -95.6719
11.675 | -95.6717
11.6753 | -95.6718
11.6757 | -95.6715
11.676 | -95.6714
11.6764 | -95.6722
11.6767 | -95.6721
11.677 | -95.6728
11.6774 | -95.6736
11.6777 | -95.6752
11.6781 | -95.6763
11.6784 | -95.6765
11.6787 | -95.6768
11.6791 | -95.6774
11.6794 | -95.6778
11.6798 | -95.6777
11.6801 | -95.6773
11.6804 | -95.6769
11.6808 | -95.677
11.6811 | -95.6769
11.6815 | -95.677
11.6818 | -95.6769
11.6821 | -95.677
11.6825 | -95.677
11.6828 | -95.677
11.6832 | -95.677
11.6835 | -95.6768
11.6838 | -95.6767
11.6842 | -95.6769
11.6845 | -95.6766
11.6849 | -95.6766
11.6852 | -95.6766
11.6856 | -95.6765
11.6859 | -95.6765
11.6862 | -95.6768
11.6866 | -95.6769
11.6869 | -95.6768
11.6873 | -95.677
11.6876 | -95.6772
11.6879 | -95.6773
11.6883 | -95.6772
11.6886 | -95.677
11.689 | -95.6771
11.6893 | -95.6777
11.6896 | -95.6774
11.69 | -95.6771
11.6903 | -95.6773
11.6907 | -95.6771
11.691 | -95.6775
11.6913 | -95.6784
11.6917 | -95.6792
11.692 | -95.6794
11.6924 | -95.6791
11.6927 | -95.6787
11.693 | -95.6787
11.6934 | -95.6791
11.6937 | -95.6796
11.6941 | -95.6796
11.6944 | -95.6795
11.6947 | -95.6794
11.6951 | -95.6793
11.6954 | -95.6796
11.6958 | -95.6797
11.6961 | -95.6797
11.6964 | -95.6798
11.6968 | -95.68
11.6971 | -95.6804
11.6975 | -95.6808
11.6978 | -95.6805
11.6982 | -95.6808
11.6985 | -95.6809
11.6988 | -95.6805
11.6992 | -95.6803
11.6995 | -95.6803
11.6999 | -95.6805
11.7002 | -95.6802
11.7005 | -95.6801
11.7009 | -95.6797
11.7012 | -95.6796
11.7016 | -95.6802
11.7019 | -95.6806
11.7022 | -95.6807
11.7026 | -95.6807
11.7029 | -95.6807
11.7033 | -95.6807
11.7036 | -95.6806
11.7039 | -95.6805
11.7043 | -95.6807
11.7046 | -95.6808
11.705 | -95.6808
11.7053 | -95.6808
11.7056 | -95.6801
11.706 | -95.6799
11.7063 | -95.6802
11.7067 | -95.6805
11.707 | -95.6805
11.7073 | -95.6805
11.7077 | -95.6805
11.708 | -95.6805
11.7084 | -95.6805
11.7087 | -95.6806
11.709 | -95.6805
11.7094 | -95.6806
11.7097 | -95.681
11.7101 | -95.6811
11.7104 | -95.6811
11.7107 | -95.681
11.7111 | -95.6808
11.7114 | -95.6812
11.7118 | -95.6813
11.7121 | -95.6816
11.7125 | -95.6815
11.7128 | -95.6824
11.7131 | -95.6824
11.7135 | -95.6825
11.7138 | -95.6822
11.7142 | -95.6817
11.7145 | -95.6817
11.7148 | -95.6816
11.7152 | -95.6816
11.7155 | -95.6814
11.7159 | -95.6809
11.7162 | -95.6809
11.7165 | -95.6814
11.7169 | -95.6814
11.7172 | -95.6814
11.7176 | -95.6815
11.7179 | -95.6812
11.7182 | -95.681
11.7186 | -95.6811
11.7189 | -95.6819
11.7193 | -95.6812
11.7196 | -95.6809
11.7199 | -95.6805
11.7203 | -95.6803
11.7206 | -95.6802
11.721 | -95.68
11.7213 | -95.6799
11.7216 | -95.6798
11.722 | -95.6796
11.7223 | -95.6798
11.7227 | -95.6802
11.723 | -95.6802
11.7233 | -95.6802
11.7237 | -95.6802
11.724 | -95.6802
11.7244 | -95.6799
11.7247 | -95.6794
11.725 | -95.6791
11.7254 | -95.6791
11.7257 | -95.6789
11.7261 | -95.6788
11.7264 | -95.6788
11.7268 | -95.679
11.7271 | -95.6793
11.7274 | -95.6795
11.7278 | -95.6797
11.7281 | -95.6795
11.7285 | -95.6792
11.7288 | -95.6793
11.7291 | -95.6797
11.7295 | -95.6796
11.7298 | -95.6795
11.7302 | -95.6794
11.7305 | -95.6793
11.7308 | -95.679
11.7312 | -95.6793
11.7315 | -95.6798
11.7319 | -95.6796
11.7322 | -95.6796
11.7325 | -95.6794
11.7329 | -95.6794
11.7332 | -95.6798
11.7336 | -95.68
11.7339 | -95.6805
11.7342 | -95.6804
11.7346 | -95.68
11.7349 | -95.6801
11.7353 | -95.6799
11.7356 | -95.6805
11.7359 | -95.6808
11.7363 | -95.681
11.7366 | -95.6811
11.737 | -95.681
11.7373 | -95.6807
11.7376 | -95.6806
11.738 | -95.6802
11.7383 | -95.6801
11.7387 | -95.6799
11.739 | -95.681
11.7393 | -95.6806
11.7397 | -95.6817
11.74 | -95.6815
11.7404 | -95.6814
11.7407 | -95.6811
11.7411 | -95.6813
11.7414 | -95.6814
11.7417 | -95.6814
11.7421 | -95.6816
11.7424 | -95.6815
11.7428 | -95.6807
11.7431 | -95.6807
11.7434 | -95.6815
11.7438 | -95.682
11.7441 | -95.6817
11.7445 | -95.6815
11.7448 | -95.681
11.7451 | -95.6805
11.7455 | -95.6804
11.7458 | -95.6803
11.7462 | -95.6804
11.7465 | -95.681
11.7468 | -95.6812
11.7472 | -95.6809
11.7475 | -95.6806
11.7479 | -95.6806
11.7482 | -95.6806
11.7485 | -95.6804
11.7489 | -95.6805
11.7492 | -95.6805
11.7496 | -95.6803
11.7499 | -95.6801
11.7502 | -95.6797
11.7506 | -95.6795
11.7509 | -95.6791
11.7513 | -95.6788
11.7516 | -95.6788
11.7519 | -95.6788
11.7523 | -95.6785
11.7526 | -95.6783
11.753 | -95.6782
11.7533 | -95.6781
11.7537 | -95.6777
11.754 | -95.6776
11.7543 | -95.6773
11.7547 | -95.677
11.755 | -95.6767
11.7554 | -95.6769
11.7557 | -95.6772
11.756 | -95.6768
11.7564 | -95.6765
11.7567 | -95.6761
11.7571 | -95.6761
11.7574 | -95.6764
11.7577 | -95.6763
11.7581 | -95.6765
11.7584 | -95.6766
11.7588 | -95.6763
11.7591 | -95.6763
11.7594 | -95.6756
11.7598 | -95.6742
11.7601 | -95.6737
11.7605 | -95.6736
11.7608 | -95.6735
11.7611 | -95.6732
11.7615 | -95.673
11.7618 | -95.6732
11.7622 | -95.6729
11.7625 | -95.6729
11.7628 | -95.6729
11.7632 | -95.6725
11.7635 | -95.6721
11.7639 | -95.6718
11.7642 | -95.6711
11.7645 | -95.6709
11.7649 | -95.6707
11.7652 | -95.671
11.7656 | -95.6708
11.7659 | -95.6705
11.7662 | -95.6705
11.7666 | -95.6709
11.7669 | -95.6709
11.7673 | -95.6709
11.7676 | -95.6708
11.768 | -95.6704
11.7683 | -95.6708
11.7686 | -95.671
11.769 | -95.671
11.7693 | -95.6709
11.7697 | -95.6705
11.77 | -95.6703
11.7703 | -95.6702
11.7707 | -95.67
11.771 | -95.6699
11.7714 | -95.6697
11.7717 | -95.6694
11.772 | -95.6691
11.7724 | -95.6685
11.7727 | -95.6683
11.7731 | -95.6683
11.7734 | -95.6681
11.7737 | -95.6678
11.7741 | -95.6672
11.7744 | -95.6664
11.7748 | -95.6658
11.7751 | -95.6656
11.7754 | -95.6655
11.7758 | -95.6651
11.7761 | -95.6641
11.7765 | -95.6638
11.7768 | -95.6634
11.7771 | -95.6619
11.7775 | -95.6605
11.7778 | -95.6603
11.7782 | -95.66
11.7785 | -95.6596
11.7788 | -95.6591
11.7792 | -95.6588
11.7795 | -95.6585
11.7799 | -95.6583
11.7802 | -95.6582
11.7805 | -95.658
11.7809 | -95.6579
11.7812 | -95.6577
11.7816 | -95.6574
11.7819 | -95.657
11.7823 | -95.6566
11.7826 | -95.6563
11.7829 | -95.6563
11.7833 | -95.6561
11.7836 | -95.6558
11.784 | -95.6556
11.7843 | -95.6555
11.7846 | -95.6553
11.785 | -95.6553
11.7853 | -95.6549
11.7857 | -95.6541
11.786 | -95.653
11.7863 | -95.6523
11.7867 | -95.6521
11.787 | -95.6516
11.7874 | -95.6516
11.7877 | -95.6515
11.788 | -95.6512
11.7884 | -95.651
11.7887 | -95.6509
11.7891 | -95.6509
11.7894 | -95.6511
11.7897 | -95.6511
11.7901 | -95.651
11.7904 | -95.6506
11.7908 | -95.6504
11.7911 | -95.6502
11.7914 | -95.65
11.7918 | -95.6496
11.7921 | -95.6494
11.7925 | -95.6491
11.7928 | -95.6487
11.7931 | -95.6487
11.7935 | -95.6489
11.7938 | -95.6488
11.7942 | -95.6489
11.7945 | -95.6487
11.7948 | -95.6486
11.7952 | -95.6486
11.7955 | -95.6484
11.7959 | -95.6481
11.7962 | -95.6479
11.7966 | -95.6477
11.7969 | -95.6475
11.7972 | -95.6478
11.7976 | -95.6478
11.7979 | -95.6474
11.7983 | -95.6472
11.7986 | -95.6469
11.7989 | -95.647
11.7993 | -95.647
11.7996 | -95.6472
11.8 | -95.647
11.8003 | -95.6468
11.8006 | -95.6468
11.801 | -95.6466
11.8013 | -95.6464
11.8017 | -95.6462
11.802 | -95.646
11.8023 | -95.6458
11.8027 | -95.6456
11.803 | -95.6456
11.8034 | -95.6458
11.8037 | -95.646
11.804 | -95.6459
11.8044 | -95.6454
11.8047 | -95.645
11.8051 | -95.645
11.8054 | -95.6453
11.8057 | -95.6451
11.8061 | -95.6446
11.8064 | -95.6441
11.8068 | -95.644
11.8071 | -95.6438
11.8074 | -95.6436
11.8078 | -95.6433
11.8081 | -95.6434
11.8085 | -95.6432
11.8088 | -95.6431
11.8092 | -95.6429
11.8095 | -95.6429
11.8098 | -95.6431
11.8102 | -95.6435
11.8105 | -95.6438
11.8109 | -95.6441
11.8112 | -95.6441
11.8115 | -95.644
11.8119 | -95.6437
11.8122 | -95.6436
11.8126 | -95.6434
11.8129 | -95.643
11.8132 | -95.6429
11.8136 | -95.6426
11.8139 | -95.642
11.8143 | -95.6419
11.8146 | -95.6418
11.8149 | -95.6415
11.8153 | -95.6415
11.8156 | -95.6414
11.816 | -95.6413
11.8163 | -95.6411
11.8166 | -95.6409
11.817 | -95.6409
11.8173 | -95.6409
11.8177 | -95.6406
11.818 | -95.6406
11.8183 | -95.6404
11.8187 | -95.6399
11.819 | -95.6397
11.8194 | -95.6402
11.8197 | -95.6399
11.82 | -95.6397
11.8204 | -95.6392
11.8207 | -95.6392
11.8211 | -95.6391
11.8214 | -95.6387
11.8217 | -95.6386
11.8221 | -95.6383
11.8224 | -95.6383
11.8228 | -95.6383
11.8231 | -95.6378
11.8234 | -95.637
11.8238 | -95.6366
11.8241 | -95.6365
11.8245 | -95.6365
11.8248 | -95.6363
11.8252 | -95.6362
11.8255 | -95.6361
11.8258 | -95.6366
11.8262 | -95.6377
11.8265 | -95.6378
11.8269 | -95.6371
11.8272 | -95.6368
11.8275 | -95.6365
11.8279 | -95.6365
11.8282 | -95.6363
11.8286 | -95.6363
11.8289 | -95.6368
11.8292 | -95.6368
11.8296 | -95.6365
11.8299 | -95.6362
11.8303 | -95.636
11.8306 | -95.6358
11.8309 | -95.6354
11.8313 | -95.6345
11.8316 | -95.6341
11.832 | -95.6342
11.8323 | -95.634
11.8326 | -95.634
11.833 | -95.6342
11.8333 | -95.6341
11.8337 | -95.6338
11.834 | -95.6336
11.8343 | -95.6335
11.8347 | -95.6335
11.835 | -95.6331
11.8354 | -95.633
11.8357 | -95.6329
11.836 | -95.6326
11.8364 | -95.6322
11.8367 | -95.6321
11.8371 | -95.6324
11.8374 | -95.6326
11.8378 | -95.6328
11.8381 | -95.633
11.8384 | -95.6329
11.8388 | -95.6331
11.8391 | -95.6332
11.8395 | -95.6333
11.8398 | -95.6334
11.8401 | -95.6333
11.8405 | -95.6331
11.8408 | -95.633
11.8412 | -95.6328
11.8415 | -95.6327
11.8418 | -95.6321
11.8422 | -95.6317
11.8425 | -95.6316
11.8429 | -95.6315
11.8432 | -95.6314
11.8435 | -95.6312
11.8439 | -95.6311
11.8442 | -95.6313
11.8446 | -95.6319
11.8449 | -95.6319
11.8452 | -95.6317
11.8456 | -95.6319
11.8459 | -95.6321
11.8463 | -95.6315
11.8466 | -95.6309
11.8469 | -95.6304
11.8473 | -95.6301
11.8476 | -95.6298
11.848 | -95.6296
11.8483 | -95.6299
11.8486 | -95.6299
11.849 | -95.6297
11.8493 | -95.6291
11.8497 | -95.6286
11.85 | -95.6282
11.8503 | -95.6279
11.8507 | -95.6275
11.851 | -95.6272
11.8514 | -95.6271
11.8517 | -95.6263
11.8521 | -95.6259
11.8524 | -95.626
11.8527 | -95.6257
11.8531 | -95.6256
11.8534 | -95.6256
11.8538 | -95.6252
11.8541 | -95.6257
11.8544 | -95.6262
11.8548 | -95.6261
11.8551 | -95.6256
11.8555 | -95.6249
11.8558 | -95.6249
11.8561 | -95.6242
11.8565 | -95.6238
11.8568 | -95.6236
11.8572 | -95.6234
11.8575 | -95.6232
11.8578 | -95.6231
11.8582 | -95.623
11.8585 | -95.623
11.8589 | -95.6229
11.8592 | -95.6233
11.8595 | -95.6229
11.8599 | -95.622
11.8602 | -95.6214
11.8606 | -95.6212
11.8609 | -95.6208
11.8612 | -95.6198
11.8616 | -95.6198
11.8619 | -95.6197
11.8623 | -95.6196
11.8626 | -95.6192
11.8629 | -95.6188
11.8633 | -95.6188
11.8636 | -95.6187
11.864 | -95.6189
11.8643 | -95.6195
11.8646 | -95.62
11.865 | -95.6198
11.8653 | -95.6196
11.8657 | -95.6194
11.866 | -95.6192
11.8664 | -95.6186
11.8667 | -95.618
11.867 | -95.6182
11.8674 | -95.6181
11.8677 | -95.6178
11.8681 | -95.6177
11.8684 | -95.6174
11.8687 | -95.6171
11.8691 | -95.6169
11.8694 | -95.6168
11.8698 | -95.6159
11.8701 | -95.6149
11.8704 | -95.6149
11.8708 | -95.6149
11.8711 | -95.6141
11.8715 | -95.6133
11.8718 | -95.6131
11.8721 | -95.6129
11.8725 | -95.6125
11.8728 | -95.6121
11.8732 | -95.6111
11.8735 | -95.6108
11.8738 | -95.6104
11.8742 | -95.6104
11.8745 | -95.6102
11.8749 | -95.6096
11.8752 | -95.6085
11.8755 | -95.6082
11.8759 | -95.6078
11.8762 | -95.6072
11.8766 | -95.607
11.8769 | -95.607
11.8772 | -95.6064
11.8776 | -95.605
11.8779 | -95.6041
11.8783 | -95.6038
11.8786 | -95.6037
11.8789 | -95.6041
11.8793 | -95.6038
11.8796 | -95.6028
11.88 | -95.6023
11.8803 | -95.6018
11.8807 | -95.6013
11.881 | -95.6018
11.8813 | -95.6012
11.8817 | -95.6012
11.882 | -95.6012
11.8824 | -95.601
11.8827 | -95.6009
11.883 | -95.601
11.8834 | -95.6004
11.8837 | -95.5991
11.8841 | -95.5984
11.8844 | -95.598
11.8847 | -95.5978
11.8851 | -95.5981
11.8854 | -95.5985
11.8858 | -95.5987
11.8861 | -95.5987
11.8864 | -95.5991
11.8868 | -95.5992
11.8871 | -95.5993
11.8875 | -95.5995
11.8878 | -95.5993
11.8881 | -95.5989
11.8885 | -95.5985
11.8888 | -95.5981
11.8892 | -95.5977
11.8895 | -95.5968
11.8898 | -95.5963
11.8902 | -95.596
11.8905 | -95.5962
11.8909 | -95.5961
11.8912 | -95.5963
11.8915 | -95.5962
11.8919 | -95.5953
11.8922 | -95.5951
11.8926 | -95.5947
11.8929 | -95.5944
11.8933 | -95.5945
11.8936 | -95.5955
11.8939 | -95.5955
11.8943 | -95.5943
11.8946 | -95.5937
11.895 | -95.5935
11.8953 | -95.5932
11.8956 | -95.593
11.896 | -95.5928
11.8963 | -95.5927
11.8967 | -95.5923
11.897 | -95.5917
11.8973 | -95.5922
11.8977 | -95.5921
11.898 | -95.5916
11.8984 | -95.5909
11.8987 | -95.5905
11.899 | -95.5909
11.8994 | -95.5909
11.8997 | -95.59
11.9001 | -95.5883
11.9004 | -95.587
11.9007 | -95.5863
11.9011 | -95.586
11.9014 | -95.586
11.9018 | -95.5862
11.9021 | -95.5864
11.9024 | -95.5866
11.9028 | -95.5864
11.9031 | -95.586
11.9035 | -95.5858
11.9038 | -95.5857
11.9041 | -95.5855
11.9045 | -95.5847
11.9048 | -95.5841
11.9052 | -95.5838
11.9055 | -95.5839
11.9058 | -95.5836
11.9062 | -95.5833
11.9065 | -95.5819
11.9069 | -95.5811
11.9072 | -95.5804
11.9076 | -95.5778
11.9079 | -95.575
11.9082 | -95.5612
11.9086 | -95.5418
11.9089 | -95.5298
11.9093 | -95.526
11.9096 | -95.5241
11.9099 | -95.5234
11.9103 | -95.5226
11.9106 | -95.5223
11.911 | -95.522
11.9113 | -95.5214
11.9116 | -95.5204
11.912 | -95.5199
11.9123 | -95.5194
11.9127 | -95.5189
11.913 | -95.5187
11.9133 | -95.5184
11.9137 | -95.5178
11.914 | -95.5175
11.9144 | -95.5166
11.9147 | -95.5159
11.915 | -95.515
11.9154 | -95.5144
11.9157 | -95.5135
11.9161 | -95.5126
11.9164 | -95.5123
11.9167 | -95.5117
11.9171 | -95.5111
11.9174 | -95.5104
11.9178 | -95.5092
11.9181 | -95.5084
11.9184 | -95.508
11.9188 | -95.5071
11.9191 | -95.5063
11.9195 | -95.5058
11.9198 | -95.5053
11.9201 | -95.5049
11.9205 | -95.5045
11.9208 | -95.5042
11.9212 | -95.504
11.9215 | -95.5033
11.9219 | -95.502
11.9222 | -95.5015
11.9225 | -95.5012
11.9229 | -95.5005
11.9232 | -95.5
11.9236 | -95.4995
11.9239 | -95.4994
11.9242 | -95.4986
11.9246 | -95.498
11.9249 | -95.4973
11.9253 | -95.4967
11.9256 | -95.4962
11.9259 | -95.4958
11.9263 | -95.4954
11.9266 | -95.4947
11.927 | -95.4937
11.9273 | -95.4923
11.9276 | -95.4915
11.928 | -95.4907
11.9283 | -95.4901
11.9287 | -95.4895
11.929 | -95.4886
11.9293 | -95.4884
11.9297 | -95.4882
11.93 | -95.4876
11.9304 | -95.4871
11.9307 | -95.4861
11.931 | -95.4855
11.9314 | -95.4849
11.9317 | -95.4841
11.9321 | -95.4838
11.9324 | -95.484
11.9327 | -95.4839
11.9331 | -95.4835
11.9334 | -95.4831
11.9338 | -95.4825
11.9341 | -95.4818
11.9344 | -95.4808
11.9348 | -95.48
11.9351 | -95.4796
11.9355 | -95.4793
11.9358 | -95.4785
11.9362 | -95.4777
11.9365 | -95.4774
11.9368 | -95.4775
11.9372 | -95.4776
11.9375 | -95.4776
11.9379 | -95.4774
11.9382 | -95.477
11.9385 | -95.4766
11.9389 | -95.4759
11.9392 | -95.4754
11.9396 | -95.4746
11.9399 | -95.4741
11.9402 | -95.4736
11.9406 | -95.4729
11.9409 | -95.4721
11.9413 | -95.4712
11.9416 | -95.4707
11.9419 | -95.4703
11.9423 | -95.4696
11.9426 | -95.469
11.943 | -95.4684
11.9433 | -95.4677
11.9436 | -95.4671
11.944 | -95.4665
11.9443 | -95.466
11.9447 | -95.4655
11.945 | -95.4648
11.9453 | -95.464
11.9457 | -95.463
11.946 | -95.4625
11.9464 | -95.4621
11.9467 | -95.4616
11.947 | -95.4608
11.9474 | -95.4601
11.9477 | -95.4597
11.9481 | -95.4591
11.9484 | -95.4584
11.9488 | -95.4578
11.9491 | -95.4572
11.9494 | -95.4568
11.9498 | -95.4564
11.9501 | -95.4558
11.9505 | -95.4553
11.9508 | -95.455
11.9511 | -95.4548
11.9515 | -95.4542
11.9518 | -95.4538
11.9522 | -95.4533
11.9525 | -95.4527
11.9528 | -95.4523
11.9532 | -95.452
11.9535 | -95.4517
11.9539 | -95.4509
11.9542 | -95.4504
11.9545 | -95.4499
11.9549 | -95.4495
11.9552 | -95.4493
11.9556 | -95.4488
11.9559 | -95.4484
11.9562 | -95.4472
11.9566 | -95.446
11.9569 | -95.4456
11.9573 | -95.4456
11.9576 | -95.445
11.9579 | -95.4445
11.9583 | -95.4439
11.9586 | -95.4436
11.959 | -95.4431
11.9593 | -95.4426
11.9596 | -95.4419
11.96 | -95.4414
11.9603 | -95.441
11.9607 | -95.4406
11.961 | -95.44
11.9613 | -95.4398
11.9617 | -95.4395
11.962 | -95.4389
11.9624 | -95.4384
11.9627 | -95.4377
11.9631 | -95.4366
11.9634 | -95.4361
11.9637 | -95.4353
11.9641 | -95.4347
11.9644 | -95.4342
11.9648 | -95.4338
11.9651 | -95.4335
11.9654 | -95.4331
11.9658 | -95.433
11.9661 | -95.432
11.9665 | -95.4313
11.9668 | -95.431
11.9671 | -95.4305
11.9675 | -95.43
11.9678 | -95.4295
11.9682 | -95.429
11.9685 | -95.4283
11.9688 | -95.4276
11.9692 | -95.427
11.9695 | -95.4262
11.9699 | -95.4257
11.9702 | -95.4252
11.9705 | -95.4247
11.9709 | -95.4244
11.9712 | -95.4239
11.9716 | -95.4232
11.9719 | -95.4226
11.9722 | -95.4221
11.9726 | -95.4213
11.9729 | -95.4208
11.9733 | -95.4202
11.9736 | -95.4193
11.9739 | -95.4186
11.9743 | -95.4181
11.9746 | -95.4175
11.975 | -95.4169
11.9753 | -95.416
11.9756 | -95.415
11.976 | -95.4141
11.9763 | -95.4136
11.9767 | -95.413
11.977 | -95.4122
11.9774 | -95.4116
11.9777 | -95.411
11.978 | -95.41
11.9784 | -95.4092
11.9787 | -95.4084
11.9791 | -95.4078
11.9794 | -95.4072
11.9797 | -95.4065
11.9801 | -95.4056
11.9804 | -95.4049
11.9808 | -95.4043
11.9811 | -95.4034
11.9814 | -95.4027
11.9818 | -95.4019
11.9821 | -95.4014
11.9825 | -95.4017
11.9828 | -95.4017
11.9831 | -95.401
11.9835 | -95.4005
11.9838 | -95.3999
11.9842 | -95.3988
11.9845 | -95.3981
11.9848 | -95.3974
11.9852 | -95.3968
11.9855 | -95.3963
11.9859 | -95.3961
11.9862 | -95.3957
11.9865 | -95.3952
11.9869 | -95.3947
11.9872 | -95.3943
11.9876 | -95.3943
11.9879 | -95.3937
11.9882 | -95.393
11.9886 | -95.3925
11.9889 | -95.392
11.9893 | -95.3913
11.9896 | -95.3903
11.9899 | -95.3895
11.9903 | -95.389
11.9906 | -95.3885
11.991 | -95.388
11.9913 | -95.3872
11.9917 | -95.3866
11.992 | -95.3859
11.9923 | -95.3862
11.9927 | -95.3864
11.993 | -95.3859
11.9934 | -95.3852
11.9937 | -95.3844
11.994 | -95.3837
11.9944 | -95.3832
11.9947 | -95.3825
11.9951 | -95.3818
11.9954 | -95.3811
11.9957 | -95.3802
11.9961 | -95.3797
11.9964 | -95.3791
11.9968 | -95.3785
11.9971 | -95.3779
11.9974 | -95.3773
11.9978 | -95.3764
11.9981 | -95.3751
11.9985 | -95.3735
11.9988 | -95.3727
11.9991 | -95.3724
11.9995 | -95.3719
11.9998 | -95.3712
12.0002 | -95.3704
12.0005 | -95.3697
12.0008 | -95.3691
12.0012 | -95.3683
12.0015 | -95.3675
12.0019 | -95.3665
12.0022 | -95.366
12.0025 | -95.3654
12.0029 | -95.3648
12.0032 | -95.3644
12.0036 | -95.3638
12.0039 | -95.3633
12.0042 | -95.3627
12.0046 | -95.3619
12.0049 | -95.3614
12.0053 | -95.361
12.0056 | -95.3611
12.006 | -95.3607
12.0063 | -95.3604
12.0066 | -95.3604
12.007 | -95.3603
12.0073 | -95.3597
12.0077 | -95.3591
12.008 | -95.3586
12.0083 | -95.3578
12.0087 | -95.3567
12.009 | -95.3561
12.0094 | -95.3557
12.0097 | -95.3551
12.01 | -95.3546
12.0104 | -95.3539
12.0107 | -95.3534
12.0111 | -95.3523
12.0114 | -95.3511
12.0117 | -95.3508
12.0121 | -95.3501
12.0124 | -95.3488
12.0128 | -95.3482
12.0131 | -95.3476
12.0134 | -95.3475
12.0138 | -95.3473
12.0141 | -95.347
12.0145 | -95.3465
12.0148 | -95.3461
12.0151 | -95.3455
12.0155 | -95.3449
12.0158 | -95.3446
12.0162 | -95.3439
12.0165 | -95.3431
12.0168 | -95.3424
12.0172 | -95.3418
12.0175 | -95.3413
12.0179 | -95.3407
12.0182 | -95.3404
12.0185 | -95.3399
12.0189 | -95.3397
12.0192 | -95.3394
12.0196 | -95.3388
12.0199 | -95.3383
12.0203 | -95.3378
12.0206 | -95.3372
12.0209 | -95.3367
12.0213 | -95.3366
12.0216 | -95.3365
12.022 | -95.3359
12.0223 | -95.3353
12.0226 | -95.3348
12.023 | -95.3341
12.0233 | -95.333
12.0237 | -95.3319
12.024 | -95.3312
12.0243 | -95.3303
12.0247 | -95.3297
12.025 | -95.3295
12.0254 | -95.3292
12.0257 | -95.3286
12.026 | -95.328
12.0264 | -95.3274
12.0267 | -95.3268
12.0271 | -95.326
12.0274 | -95.3256
12.0277 | -95.3253
12.0281 | -95.325
12.0284 | -95.3243
12.0288 | -95.3235
12.0291 | -95.3228
12.0294 | -95.3221
12.0298 | -95.3214
12.0301 | -95.3209
12.0305 | -95.3203
12.0308 | -95.3194
12.0311 | -95.319
12.0315 | -95.318
12.0318 | -95.3171
12.0322 | -95.3165
12.0325 | -95.3158
12.0328 | -95.3155
12.0332 | -95.3154
12.0335 | -95.3147
12.0339 | -95.314
12.0342 | -95.3135
12.0346 | -95.313
12.0349 | -95.3124
12.0352 | -95.3118
12.0356 | -95.3111
12.0359 | -95.3104
12.0363 | -95.3098
12.0366 | -95.3092
12.0369 | -95.3086
12.0373 | -95.3081
12.0376 | -95.3074
12.038 | -95.3068
12.0383 | -95.306
12.0386 | -95.3049
12.039 | -95.3042
12.0393 | -95.3036
12.0397 | -95.3029
12.04 | -95.3025
12.0403 | -95.3018
12.0407 | -95.3006
12.041 | -95.2995
12.0414 | -95.2986
12.0417 | -95.298
12.042 | -95.2974
12.0424 | -95.2967
12.0427 | -95.2959
12.0431 | -95.2952
12.0434 | -95.2945
12.0437 | -95.294
12.0441 | -95.2934
12.0444 | -95.293
12.0448 | -95.2932
12.0451 | -95.2934
12.0454 | -95.2927
12.0458 | -95.2922
12.0461 | -95.2915
12.0465 | -95.2908
12.0468 | -95.2901
12.0472 | -95.2894
12.0475 | -95.2885
12.0478 | -95.2879
12.0482 | -95.2874
12.0485 | -95.2869
12.0489 | -95.2866
12.0492 | -95.2861
12.0495 | -95.2856
12.0499 | -95.285
12.0502 | -95.2845
12.0506 | -95.2837
12.0509 | -95.2828
12.0512 | -95.2816
12.0516 | -95.2806
12.0519 | -95.28
12.0523 | -95.2795
12.0526 | -95.279
12.0529 | -95.2789
12.0533 | -95.2783
12.0536 | -95.2779
12.054 | -95.2773
12.0543 | -95.2772
12.0546 | -95.2769
12.055 | -95.2762
12.0553 | -95.2755
12.0557 | -95.2748
12.056 | -95.2743
12.0563 | -95.2737
12.0567 | -95.2731
12.057 | -95.2723
12.0574 | -95.2718
12.0577 | -95.2714
12.058 | -95.2708
12.0584 | -95.2698
12.0587 | -95.269
12.0591 | -95.2682
12.0594 | -95.2674
12.0597 | -95.2667
12.0601 | -95.2661
12.0604 | -95.2654
12.0608 | -95.2647
12.0611 | -95.2642
12.0615 | -95.2635
12.0618 | -95.2627
12.0621 | -95.2623
12.0625 | -95.2618
12.0628 | -95.2611
12.0632 | -95.2607
12.0635 | -95.26
12.0638 | -95.2595
12.0642 | -95.2588
12.0645 | -95.2579
12.0649 | -95.2571
12.0652 | -95.2566
12.0655 | -95.256
12.0659 | -95.2556
12.0662 | -95.255
12.0666 | -95.2541
12.0669 | -95.2532
12.0672 | -95.2524
12.0676 | -95.2517
12.0679 | -95.251
12.0683 | -95.2501
12.0686 | -95.2493
12.0689 | -95.2486
12.0693 | -95.2481
12.0696 | -95.2478
12.07 | -95.2473
12.0703 | -95.2468
12.0706 | -95.2462
12.071 | -95.2452
12.0713 | -95.2442
12.0717 | -95.2438
12.072 | -95.2435
12.0723 | -95.2427
12.0727 | -95.2419
12.073 | -95.241
12.0734 | -95.2405
12.0737 | -95.2398
12.074 | -95.2392
12.0744 | -95.2387
12.0747 | -95.238
12.0751 | -95.2372
12.0754 | -95.236
12.0758 | -95.2351
12.0761 | -95.2344
12.0764 | -95.2338
12.0768 | -95.2333
12.0771 | -95.2325
12.0775 | -95.2317
12.0778 | -95.2308
12.0781 | -95.2301
12.0785 | -95.2295
12.0788 | -95.2287
12.0792 | -95.2279
12.0795 | -95.2277
12.0798 | -95.2269
12.0802 | -95.2263
12.0805 | -95.2258
12.0809 | -95.2249
12.0812 | -95.2241
12.0815 | -95.2232
12.0819 | -95.2209
12.0822 | -95.2193
12.0826 | -95.2157
12.0829 | -95.2119
12.0832 | -95.2098
12.0836 | -95.209
12.0839 | -95.2081
12.0843 | -95.2059
12.0846 | -95.2041
12.0849 | -95.2023
12.0853 | -95.2002
12.0856 | -95.1981
12.086 | -95.1973
12.0863 | -95.1966
12.0866 | -95.1955
12.087 | -95.1945
12.0873 | -95.1922
12.0877 | -95.1895
12.088 | -95.1857
12.0884 | -95.1825
12.0887 | -95.182
12.089 | -95.1817
12.0894 | -95.1802
12.0897 | -95.1784
12.0901 | -95.177
12.0904 | -95.1758
12.0907 | -95.1746
12.0911 | -95.1738
12.0914 | -95.1729
12.0918 | -95.1717
12.0921 | -95.1704
12.0924 | -95.1691
12.0928 | -95.168
12.0931 | -95.167
12.0935 | -95.1664
12.0938 | -95.1656
12.0941 | -95.1647
12.0945 | -95.1636
12.0948 | -95.1622
12.0952 | -95.1608
12.0955 | -95.1597
12.0958 | -95.1588
12.0962 | -95.1579
12.0965 | -95.1568
12.0969 | -95.1561
12.0972 | -95.1557
12.0975 | -95.1547
12.0979 | -95.1535
12.0982 | -95.1525
12.0986 | -95.1516
12.0989 | -95.1507
12.0992 | -95.1498
12.0996 | -95.149
12.0999 | -95.1482
12.1003 | -95.1473
12.1006 | -95.1465
12.1009 | -95.1458
12.1013 | -95.145
12.1016 | -95.144
12.102 | -95.1432
12.1023 | -95.1424
12.1027 | -95.1418
12.103 | -95.1402
12.1033 | -95.1394
12.1037 | -95.1385
12.104 | -95.1376
12.1044 | -95.1366
12.1047 | -95.1359
12.105 | -95.1353
12.1054 | -95.1344
12.1057 | -95.1337
12.1061 | -95.1328
12.1064 | -95.1322
12.1067 | -95.1314
12.1071 | -95.1307
12.1074 | -95.1303
12.1078 | -95.1298
12.1081 | -95.1289
12.1084 | -95.1279
12.1088 | -95.1272
12.1091 | -95.1263
12.1095 | -95.1251
12.1098 | -95.1237
12.1101 | -95.1225
12.1105 | -95.1215
12.1108 | -95.1206
12.1112 | -95.1197
12.1115 | -95.1188
12.1118 | -95.1178
12.1122 | -95.117
12.1125 | -95.1161
12.1129 | -95.1151
12.1132 | -95.1143
12.1135 | -95.1132
12.1139 | -95.1123
12.1142 | -95.1115
12.1146 | -95.1109
12.1149 | -95.1106
12.1152 | -95.1096
12.1156 | -95.1091
12.1159 | -95.1088
12.1163 | -95.108
12.1166 | -95.1071
12.117 | -95.1063
12.1173 | -95.1055
12.1176 | -95.1046
12.118 | -95.104
12.1183 | -95.1031
12.1187 | -95.1021
12.119 | -95.1013
12.1193 | -95.1008
12.1197 | -95.0999
12.12 | -95.099
12.1204 | -95.0982
12.1207 | -95.0976
12.121 | -95.0973
12.1214 | -95.0966
12.1217 | -95.0957
12.1221 | -95.0948
12.1224 | -95.094
12.1227 | -95.0931
12.1231 | -95.0923
12.1234 | -95.0914
12.1238 | -95.0908
12.1241 | -95.0899
12.1244 | -95.0888
12.1248 | -95.0878
12.1251 | -95.087
12.1255 | -95.0863
12.1258 | -95.0855
12.1261 | -95.0846
12.1265 | -95.0837
12.1268 | -95.0828
12.1272 | -95.0817
12.1275 | -95.0805
12.1278 | -95.0795
12.1282 | -95.0784
12.1285 | -95.0776
12.1289 | -95.0764
12.1292 | -95.0754
12.1295 | -95.0744
12.1299 | -95.0734
12.1302 | -95.0726
12.1306 | -95.0717
12.1309 | -95.0705
12.1313 | -95.0695
12.1316 | -95.0686
12.1319 | -95.0676
12.1323 | -95.0668
12.1326 | -95.0658
12.133 | -95.0644
12.1333 | -95.0632
12.1336 | -95.0626
12.134 | -95.0619
12.1343 | -95.0611
12.1347 | -95.0603
12.135 | -95.0596
12.1353 | -95.0585
12.1357 | -95.0575
12.136 | -95.0566
12.1364 | -95.0558
12.1367 | -95.0548
12.137 | -95.0539
12.1374 | -95.0531
12.1377 | -95.0523
12.1381 | -95.0514
12.1384 | -95.0505
12.1387 | -95.0495
12.1391 | -95.0486
12.1394 | -95.0477
12.1398 | -95.0469
12.1401 | -95.0462
12.1404 | -95.046
12.1408 | -95.0455
12.1411 | -95.0447
12.1415 | -95.0439
12.1418 | -95.0429
12.1421 | -95.0413
12.1425 | -95.0402
12.1428 | -95.0399
12.1432 | -95.0391
12.1435 | -95.0383
12.1438 | -95.037
12.1442 | -95.0362
12.1445 | -95.0355
12.1449 | -95.035
12.1452 | -95.0341
12.1456 | -95.0331
12.1459 | -95.032
12.1462 | -95.0312
12.1466 | -95.0303
12.1469 | -95.0293
12.1473 | -95.0285
12.1476 | -95.0279
12.1479 | -95.0275
12.1483 | -95.0269
12.1486 | -95.026
12.149 | -95.0255
12.1493 | -95.0245
12.1496 | -95.0236
12.15 | -95.0228
12.1503 | -95.0217
12.1507 | -95.0207
12.151 | -95.0195
12.1513 | -95.0186
12.1517 | -95.018
12.152 | -95.0172
12.1524 | -95.0162
12.1527 | -95.015
12.153 | -95.0139
12.1534 | -95.0132
12.1537 | -95.0125
12.1541 | -95.0121
12.1544 | -95.0113
12.1547 | -95.0101
12.1551 | -95.0089
12.1554 | -95.0078
12.1558 | -95.0069
12.1561 | -95.0065
12.1564 | -95.0061
12.1568 | -95.0052
12.1571 | -95.0043
12.1575 | -95.0034
12.1578 | -95.0025
12.1581 | -95.0018
12.1585 | -95.0013
12.1588 | -95.0007
12.1592 | -94.9997
12.1595 | -94.9985
12.1599 | -94.9976
12.1602 | -94.9968
12.1605 | -94.9959
12.1609 | -94.995
12.1612 | -94.9942
12.1616 | -94.9937
12.1619 | -94.9931
12.1622 | -94.9924
12.1626 | -94.992
12.1629 | -94.9912
12.1633 | -94.9904
12.1636 | -94.9903
12.1639 | -94.9899
12.1643 | -94.989
12.1646 | -94.9881
12.165 | -94.9876
12.1653 | -94.9868
12.1656 | -94.9861
12.166 | -94.9855
12.1663 | -94.9851
12.1667 | -94.9844
12.167 | -94.9837
12.1673 | -94.9831
12.1677 | -94.9824
12.168 | -94.9818
12.1684 | -94.9811
12.1687 | -94.9804
12.169 | -94.9796
12.1694 | -94.979
12.1697 | -94.9782
12.1701 | -94.9771
12.1704 | -94.9762
12.1707 | -94.9754
12.1711 | -94.9744
12.1714 | -94.9735
12.1718 | -94.973
12.1721 | -94.9723
12.1724 | -94.9717
12.1728 | -94.9711
12.1731 | -94.9704
12.1735 | -94.9696
12.1738 | -94.969
12.1742 | -94.9684
12.1745 | -94.9674
12.1748 | -94.9665
12.1752 | -94.9657
12.1755 | -94.9648
12.1759 | -94.9639
12.1762 | -94.9631
12.1765 | -94.9623
12.1769 | -94.9614
12.1772 | -94.9606
12.1776 | -94.9598
12.1779 | -94.9589
12.1782 | -94.958
12.1786 | -94.9572
12.1789 | -94.9563
12.1793 | -94.9553
12.1796 | -94.9545
12.1799 | -94.9537
12.1803 | -94.9527
12.1806 | -94.9518
12.181 | -94.9512
12.1813 | -94.9506
12.1816 | -94.9501
12.182 | -94.9495
12.1823 | -94.9487
12.1827 | -94.9481
12.183 | -94.9477
12.1833 | -94.9468
12.1837 | -94.9461
12.184 | -94.945
12.1844 | -94.9443
12.1847 | -94.9436
12.185 | -94.9429
12.1854 | -94.9419
12.1857 | -94.9409
12.1861 | -94.9399
12.1864 | -94.9392
12.1868 | -94.9385
12.1871 | -94.9376
12.1874 | -94.9368
12.1878 | -94.9363
12.1881 | -94.9357
12.1885 | -94.9347
12.1888 | -94.9339
12.1891 | -94.933
12.1895 | -94.9324
12.1898 | -94.9317
12.1902 | -94.9308
12.1905 | -94.93
12.1908 | -94.9289
12.1912 | -94.9281
12.1915 | -94.9269
12.1919 | -94.9259
12.1922 | -94.9249
12.1925 | -94.924
12.1929 | -94.9231
12.1932 | -94.9223
12.1936 | -94.9215
12.1939 | -94.9209
12.1942 | -94.92
12.1946 | -94.9196
12.1949 | -94.9191
12.1953 | -94.9183
12.1956 | -94.9176
12.1959 | -94.917
12.1963 | -94.9164
12.1966 | -94.9156
12.197 | -94.915
12.1973 | -94.9147
12.1976 | -94.9138
12.198 | -94.9127
12.1983 | -94.9116
12.1987 | -94.9106
12.199 | -94.91
12.1993 | -94.9093
12.1997 | -94.9085
12.2 | -94.9076
12.2004 | -94.9068
12.2007 | -94.9059
12.2011 | -94.9057
12.2014 | -94.905
12.2017 | -94.9043
12.2021 | -94.9039
12.2024 | -94.9035
12.2028 | -94.9027
12.2031 | -94.902
12.2034 | -94.9013
12.2038 | -94.9007
12.2041 | -94.9
12.2045 | -94.899
12.2048 | -94.8982
12.2051 | -94.8976
12.2055 | -94.8969
12.2058 | -94.8962
12.2062 | -94.8954
12.2065 | -94.8946
12.2068 | -94.894
12.2072 | -94.8933
12.2075 | -94.8926
12.2079 | -94.8917
12.2082 | -94.8907
12.2085 | -94.8901
12.2089 | -94.8892
12.2092 | -94.8886
12.2096 | -94.888
12.2099 | -94.8874
12.2102 | -94.8868
12.2106 | -94.886
12.2109 | -94.8857
12.2113 | -94.8851
12.2116 | -94.8844
12.2119 | -94.8836
12.2123 | -94.8829
12.2126 | -94.8823
12.213 | -94.8817
12.2133 | -94.8811
12.2136 | -94.8802
12.214 | -94.8796
12.2143 | -94.879
12.2147 | -94.8783
12.215 | -94.8776
12.2154 | -94.8769
12.2157 | -94.8763
12.216 | -94.8757
12.2164 | -94.875
12.2167 | -94.8742
12.2171 | -94.8735
12.2174 | -94.8729
12.2177 | -94.8722
12.2181 | -94.8715
12.2184 | -94.8709
12.2188 | -94.8704
12.2191 | -94.87
12.2194 | -94.8695
12.2198 | -94.8688
12.2201 | -94.8682
12.2205 | -94.8673
12.2208 | -94.867
12.2211 | -94.8664
12.2215 | -94.8659
12.2218 | -94.8652
12.2222 | -94.8644
12.2225 | -94.8637
12.2228 | -94.8632
12.2232 | -94.8626
12.2235 | -94.8621
12.2239 | -94.8616
12.2242 | -94.861
12.2245 | -94.8605
12.2249 | -94.8601
12.2252 | -94.8595
12.2256 | -94.859
12.2259 | -94.8586
12.2262 | -94.8579
12.2266 | -94.8574
12.2269 | -94.8568
12.2273 | -94.8562
12.2276 | -94.8556
12.2279 | -94.8551
12.2283 | -94.8545
12.2286 | -94.8541
12.229 | -94.8537
12.2293 | -94.8535
12.2297 | -94.8532
12.23 | -94.8528
12.2303 | -94.8523
12.2307 | -94.8523
12.231 | -94.8518
12.2314 | -94.8513
12.2317 | -94.8513
12.232 | -94.8509
12.2324 | -94.8504
12.2327 | -94.8501
12.2331 | -94.8495
12.2334 | -94.8488
12.2337 | -94.8482
12.2341 | -94.8476
12.2344 | -94.8471
12.2348 | -94.8467
12.2351 | -94.8462
12.2354 | -94.8464
12.2358 | -94.8463
12.2361 | -94.8458
12.2365 | -94.8455
12.2368 | -94.8452
12.2371 | -94.8447
12.2375 | -94.8441
12.2378 | -94.8436
12.2382 | -94.843
12.2385 | -94.8426
12.2388 | -94.842
12.2392 | -94.8412
12.2395 | -94.8405
12.2399 | -94.8402
12.2402 | -94.84
12.2405 | -94.84
12.2409 | -94.8396
12.2412 | -94.8392
12.2416 | -94.8387
12.2419 | -94.8381
12.2423 | -94.8375
12.2426 | -94.8369
12.2429 | -94.8364
12.2433 | -94.8359
12.2436 | -94.8358
12.244 | -94.8355
12.2443 | -94.8349
12.2446 | -94.8343
12.245 | -94.8338
12.2453 | -94.8333
12.2457 | -94.8328
12.246 | -94.8323
12.2463 | -94.8318
12.2467 | -94.8312
12.247 | -94.8305
12.2474 | -94.83
12.2477 | -94.8295
12.248 | -94.8291
12.2484 | -94.8286
12.2487 | -94.8279
12.2491 | -94.8274
12.2494 | -94.8268
12.2497 | -94.8262
12.2501 | -94.8259
12.2504 | -94.8261
12.2508 | -94.826
12.2511 | -94.8256
12.2514 | -94.8253
12.2518 | -94.8246
12.2521 | -94.8236
12.2525 | -94.8231
12.2528 | -94.8229
12.2531 | -94.8225
12.2535 | -94.8217
12.2538 | -94.821
12.2542 | -94.8204
12.2545 | -94.8199
12.2548 | -94.8194
12.2552 | -94.8189
12.2555 | -94.8184
12.2559 | -94.8181
12.2562 | -94.8177
12.2566 | -94.8174
12.2569 | -94.8174
12.2572 | -94.817
12.2576 | -94.8165
12.2579 | -94.8161
12.2583 | -94.8154
12.2586 | -94.8147
12.2589 | -94.8142
12.2593 | -94.8137
12.2596 | -94.813
12.26 | -94.8133
12.2603 | -94.813
12.2606 | -94.813
12.261 | -94.8132
12.2613 | -94.8131
12.2617 | -94.8129
12.262 | -94.8125
12.2623 | -94.8124
12.2627 | -94.8121
12.263 | -94.8116
12.2634 | -94.8111
12.2637 | -94.8106
12.264 | -94.8105
12.2644 | -94.8104
12.2647 | -94.8101
12.2651 | -94.8097
12.2654 | -94.809
12.2657 | -94.8085
12.2661 | -94.8081
12.2664 | -94.8078
12.2668 | -94.8074
12.2671 | -94.8069
12.2674 | -94.8065
12.2678 | -94.8061
12.2681 | -94.806
12.2685 | -94.806
12.2688 | -94.8057
12.2691 | -94.8053
12.2695 | -94.805
12.2698 | -94.8055
12.2702 | -94.8052
12.2705 | -94.8049
12.2709 | -94.8045
12.2712 | -94.8038
12.2715 | -94.8033
12.2719 | -94.8031
12.2722 | -94.803
12.2726 | -94.8026
12.2729 | -94.802
12.2732 | -94.8013
12.2736 | -94.8007
12.2739 | -94.8002
12.2743 | -94.7999
12.2746 | -94.7997
12.2749 | -94.7995
12.2753 | -94.7992
12.2756 | -94.7988
12.276 | -94.7984
12.2763 | -94.7982
12.2766 | -94.7979
12.277 | -94.7977
12.2773 | -94.7976
12.2777 | -94.7974
12.278 | -94.7971
12.2783 | -94.7969
12.2787 | -94.7968
12.279 | -94.7967
12.2794 | -94.7965
12.2797 | -94.7966
12.28 | -94.7967
12.2804 | -94.7965
12.2807 | -94.7962
12.2811 | -94.7955
12.2814 | -94.795
12.2817 | -94.7949
12.2821 | -94.7949
12.2824 | -94.7948
12.2828 | -94.7948
12.2831 | -94.7947
12.2835 | -94.7944
12.2838 | -94.7941
12.2841 | -94.7944
12.2845 | -94.7946
12.2848 | -94.7945
12.2852 | -94.7944
12.2855 | -94.7941
12.2858 | -94.7938
12.2862 | -94.7934
12.2865 | -94.7932
12.2869 | -94.7931
12.2872 | -94.7929
12.2875 | -94.7925
12.2879 | -94.7923
12.2882 | -94.7921
12.2886 | -94.7916
12.2889 | -94.7913
12.2892 | -94.7911
12.2896 | -94.7919
12.2899 | -94.7918
12.2903 | -94.7917
12.2906 | -94.7916
12.2909 | -94.7914
12.2913 | -94.7913
12.2916 | -94.7914
12.292 | -94.7916
12.2923 | -94.7913
12.2926 | -94.791
12.293 | -94.7908
12.2933 | -94.7908
12.2937 | -94.7907
12.294 | -94.7909
12.2943 | -94.7909
12.2947 | -94.7908
12.295 | -94.7906
12.2954 | -94.7908
12.2957 | -94.7908
12.296 | -94.7908
12.2964 | -94.7909
12.2967 | -94.7908
12.2971 | -94.7905
12.2974 | -94.7902
12.2978 | -94.7901
12.2981 | -94.7902
12.2984 | -94.7905
12.2988 | -94.7909
12.2991 | -94.7912
12.2995 | -94.791
12.2998 | -94.7911
12.3001 | -94.791
12.3005 | -94.7911
12.3008 | -94.7913
12.3012 | -94.7914
12.3015 | -94.7912
12.3018 | -94.7915
12.3022 | -94.7922
12.3025 | -94.7926
12.3029 | -94.7925
12.3032 | -94.7925
12.3035 | -94.7927
12.3039 | -94.793
12.3042 | -94.7929
12.3046 | -94.7929
12.3049 | -94.7929
12.3052 | -94.7927
12.3056 | -94.7926
12.3059 | -94.7928
12.3063 | -94.7926
12.3066 | -94.7925
12.3069 | -94.7927
12.3073 | -94.7931
12.3076 | -94.7933
12.308 | -94.7931
12.3083 | -94.7928
12.3086 | -94.7931
12.309 | -94.7934
12.3093 | -94.7943
12.3097 | -94.7949
12.31 | -94.7953
12.3103 | -94.7952
12.3107 | -94.7953
12.311 | -94.7952
12.3114 | -94.7951
12.3117 | -94.7948
12.3121 | -94.7947
12.3124 | -94.795
12.3127 | -94.795
12.3131 | -94.7949
12.3134 | -94.7948
12.3138 | -94.7946
12.3141 | -94.7946
12.3144 | -94.7947
12.3148 | -94.795
12.3151 | -94.7952
12.3155 | -94.7952
12.3158 | -94.7952
12.3161 | -94.7952
12.3165 | -94.7951
12.3168 | -94.7952
12.3172 | -94.7953
12.3175 | -94.7954
12.3178 | -94.795
12.3182 | -94.7945
12.3185 | -94.794
12.3189 | -94.7939
12.3192 | -94.7936
12.3195 | -94.7933
12.3199 | -94.7929
12.3202 | -94.7925
12.3206 | -94.7925
12.3209 | -94.7927
12.3212 | -94.7927
12.3216 | -94.7927
12.3219 | -94.7934
12.3223 | -94.7938
12.3226 | -94.7942
12.3229 | -94.7946
12.3233 | -94.7949
12.3236 | -94.795
12.324 | -94.7952
12.3243 | -94.7953
12.3246 | -94.7956
12.325 | -94.7958
12.3253 | -94.7959
12.3257 | -94.796
12.326 | -94.7962
12.3264 | -94.7966
12.3267 | -94.7971
12.327 | -94.797
12.3274 | -94.7971
12.3277 | -94.7973
12.3281 | -94.7977
12.3284 | -94.7978
12.3287 | -94.7982
12.3291 | -94.7986
12.3294 | -94.7992
12.3298 | -94.7998
12.3301 | -94.8004
12.3304 | -94.8006
12.3308 | -94.8008
12.3311 | -94.801
12.3315 | -94.8012
12.3318 | -94.8013
12.3321 | -94.8012
12.3325 | -94.8013
12.3328 | -94.8017
12.3332 | -94.8026
12.3335 | -94.8029
12.3338 | -94.8029
12.3342 | -94.8031
12.3345 | -94.8033
12.3349 | -94.8034
12.3352 | -94.8035
12.3355 | -94.8036
12.3359 | -94.804
12.3362 | -94.8044
12.3366 | -94.8045
12.3369 | -94.8046
12.3372 | -94.8048
12.3376 | -94.8053
12.3379 | -94.8056
12.3383 | -94.8069
12.3386 | -94.8074
12.3389 | -94.8081
12.3393 | -94.8088
12.3396 | -94.8092
12.34 | -94.8095
12.3403 | -94.81
12.3407 | -94.8105
12.341 | -94.8111
12.3413 | -94.8113
12.3417 | -94.8115
12.342 | -94.8119
12.3424 | -94.8123
12.3427 | -94.8125
12.343 | -94.8127
12.3434 | -94.8129
12.3437 | -94.813
12.3441 | -94.8132
12.3444 | -94.8135
12.3447 | -94.8137
12.3451 | -94.8141
12.3454 | -94.8145
12.3458 | -94.8147
12.3461 | -94.8159
12.3464 | -94.8173
12.3468 | -94.8178
12.3471 | -94.818
12.3475 | -94.8182
12.3478 | -94.8185
12.3481 | -94.8188
12.3485 | -94.8198
12.3488 | -94.8203
12.3492 | -94.8209
12.3495 | -94.8216
12.3498 | -94.8221
12.3502 | -94.8226
12.3505 | -94.8231
12.3509 | -94.8235
12.3512 | -94.8238
12.3515 | -94.8243
12.3519 | -94.8247
12.3522 | -94.8254
12.3526 | -94.826
12.3529 | -94.8261
12.3532 | -94.8263
12.3536 | -94.8269
12.3539 | -94.8272
12.3543 | -94.827
12.3546 | -94.8272
12.355 | -94.8274
12.3553 | -94.8277
12.3556 | -94.8281
12.356 | -94.8287
12.3563 | -94.8294
12.3567 | -94.8299
12.357 | -94.8296
12.3573 | -94.8287
12.3577 | -94.8288
12.358 | -94.8292
12.3584 | -94.8302
12.3587 | -94.8306
12.359 | -94.8313
12.3594 | -94.8318
12.3597 | -94.8322
12.3601 | -94.8327
12.3604 | -94.8331
12.3607 | -94.8336
12.3611 | -94.8341
12.3614 | -94.8347
12.3618 | -94.835
12.3621 | -94.8351
12.3624 | -94.8358
12.3628 | -94.8364
12.3631 | -94.8367
12.3635 | -94.8371
12.3638 | -94.8375
12.3641 | -94.838
12.3645 | -94.8383
12.3648 | -94.8386
12.3652 | -94.8394
12.3655 | -94.8402
12.3658 | -94.8407
12.3662 | -94.841
12.3665 | -94.8415
12.3669 | -94.8418
12.3672 | -94.8423
12.3675 | -94.8428
12.3679 | -94.8434
12.3682 | -94.8448
12.3686 | -94.8453
12.3689 | -94.8462
12.3693 | -94.8472
12.3696 | -94.8482
12.3699 | -94.8488
12.3703 | -94.8494
12.3706 | -94.85
12.371 | -94.8505
12.3713 | -94.851
12.3716 | -94.8515
12.372 | -94.852
12.3723 | -94.8525
12.3727 | -94.8528
12.373 | -94.8537
12.3733 | -94.8542
12.3737 | -94.8545
12.374 | -94.8549
12.3744 | -94.856
12.3747 | -94.8575
12.375 | -94.8586
12.3754 | -94.8595
12.3757 | -94.8602
12.3761 | -94.8606
12.3764 | -94.8609
12.3767 | -94.8614
12.3771 | -94.862
12.3774 | -94.8626
12.3778 | -94.8632
12.3781 | -94.8648
12.3784 | -94.8657
12.3788 | -94.8661
12.3791 | -94.8667
12.3795 | -94.8674
12.3798 | -94.8681
12.3801 | -94.8688
12.3805 | -94.8694
12.3808 | -94.8698
12.3812 | -94.8703
12.3815 | -94.8712
12.3819 | -94.872
12.3822 | -94.8725
12.3825 | -94.8731
12.3829 | -94.8739
12.3832 | -94.8746
12.3836 | -94.8753
12.3839 | -94.8761
12.3842 | -94.8771
12.3846 | -94.8777
12.3849 | -94.8787
12.3853 | -94.8798
12.3856 | -94.8805
12.3859 | -94.8812
12.3863 | -94.8818
12.3866 | -94.8827
12.387 | -94.8838
12.3873 | -94.8845
12.3876 | -94.8861
12.388 | -94.887
12.3883 | -94.8877
12.3887 | -94.8884
12.389 | -94.8894
12.3893 | -94.8904
12.3897 | -94.8911
12.39 | -94.8919
12.3904 | -94.8927
12.3907 | -94.8935
12.391 | -94.8941
12.3914 | -94.8946
12.3917 | -94.8951
12.3921 | -94.8962
12.3924 | -94.8969
12.3927 | -94.8975
12.3931 | -94.898
12.3934 | -94.8988
12.3938 | -94.8996
12.3941 | -94.9003
12.3944 | -94.9012
12.3948 | -94.9021
12.3951 | -94.9031
12.3955 | -94.9036
12.3958 | -94.9043
12.3962 | -94.905
12.3965 | -94.906
12.3968 | -94.9073
12.3972 | -94.9082
12.3975 | -94.9099
12.3979 | -94.9109
12.3982 | -94.9121
12.3985 | -94.9131
12.3989 | -94.9141
12.3992 | -94.9152
12.3996 | -94.916
12.3999 | -94.9168
12.4002 | -94.9177
12.4006 | -94.9191
12.4009 | -94.9204
12.4013 | -94.9212
12.4016 | -94.9222
12.4019 | -94.9238
12.4023 | -94.925
12.4026 | -94.9257
12.403 | -94.9266
12.4033 | -94.9283
12.4036 | -94.9295
12.404 | -94.9305
12.4043 | -94.9313
12.4047 | -94.9321
12.405 | -94.933
12.4053 | -94.934
12.4057 | -94.9344
12.406 | -94.9352
12.4064 | -94.9364
12.4067 | -94.9374
12.407 | -94.9383
12.4074 | -94.9401
12.4077 | -94.9411
12.4081 | -94.9419
12.4084 | -94.9428
12.4087 | -94.9438
12.4091 | -94.9445
12.4094 | -94.9457
12.4098 | -94.9466
12.4101 | -94.9474
12.4105 | -94.948
12.4108 | -94.9487
12.4111 | -94.9499
12.4115 | -94.9511
12.4118 | -94.9522
12.4122 | -94.9531
12.4125 | -94.9539
12.4128 | -94.9551
12.4132 | -94.956
12.4135 | -94.9569
12.4139 | -94.9578
12.4142 | -94.9589
12.4145 | -94.9601
12.4149 | -94.9611
12.4152 | -94.9622
12.4156 | -94.9631
12.4159 | -94.9641
12.4162 | -94.965
12.4166 | -94.9658
12.4169 | -94.9665
12.4173 | -94.968
12.4176 | -94.9691
12.4179 | -94.9698
12.4183 | -94.9705
12.4186 | -94.9714
12.419 | -94.9724
12.4193 | -94.9734
12.4196 | -94.9742
12.42 | -94.9751
12.4203 | -94.976
12.4207 | -94.977
12.421 | -94.9778
12.4213 | -94.9787
12.4217 | -94.9796
12.422 | -94.9805
12.4224 | -94.9813
12.4227 | -94.9822
12.423 | -94.9833
12.4234 | -94.9845
12.4237 | -94.9857
12.4241 | -94.9867
12.4244 | -94.9878
12.4248 | -94.9887
12.4251 | -94.9896
12.4254 | -94.9907
12.4258 | -94.9916
12.4261 | -94.9924
12.4265 | -94.9933
12.4268 | -94.9943
12.4271 | -94.9959
12.4275 | -94.9967
12.4278 | -94.9977
12.4282 | -94.9988
12.4285 | -94.9999
12.4288 | -95.001
12.4292 | -95.0021
12.4295 | -95.0032
12.4299 | -95.0044
12.4302 | -95.0055
12.4305 | -95.0069
12.4309 | -95.0085
12.4312 | -95.0101
12.4316 | -95.0112
12.4319 | -95.0123
12.4322 | -95.0137
12.4326 | -95.015
12.4329 | -95.0163
12.4333 | -95.0174
12.4336 | -95.0185
12.4339 | -95.0197
12.4343 | -95.0206
12.4346 | -95.0214
12.435 | -95.0221
12.4353 | -95.0232
12.4356 | -95.024
12.436 | -95.025
12.4363 | -95.0261
12.4367 | -95.0273
12.437 | -95.0289
12.4374 | -95.0302
12.4377 | -95.0315
12.438 | -95.0329
12.4384 | -95.0342
12.4387 | -95.0353
12.4391 | -95.0364
12.4394 | -95.0374
12.4397 | -95.0382
12.4401 | -95.04
12.4404 | -95.0416
12.4408 | -95.0427
12.4411 | -95.0438
12.4414 | -95.0448
12.4418 | -95.0461
12.4421 | -95.0477
12.4425 | -95.0493
12.4428 | -95.0508
12.4431 | -95.052
12.4435 | -95.0533
12.4438 | -95.0549
12.4442 | -95.0566
12.4445 | -95.058
12.4448 | -95.0591
12.4452 | -95.0602
12.4455 | -95.0615
12.4459 | -95.0628
12.4462 | -95.0638
12.4465 | -95.0651
12.4469 | -95.0666
12.4472 | -95.068
12.4476 | -95.0696
12.4479 | -95.071
12.4482 | -95.0721
12.4486 | -95.0729
12.4489 | -95.074
12.4493 | -95.075
12.4496 | -95.0762
12.4499 | -95.0777
12.4503 | -95.0799
12.4506 | -95.0812
12.451 | -95.0825
12.4513 | -95.0841
12.4517 | -95.0855
12.452 | -95.0867
12.4523 | -95.088
12.4527 | -95.0894
12.453 | -95.0905
12.4534 | -95.0915
12.4537 | -95.0927
12.454 | -95.094
12.4544 | -95.0953
12.4547 | -95.0965
12.4551 | -95.0979
12.4554 | -95.0995
12.4557 | -95.1007
12.4561 | -95.1018
12.4564 | -95.1036
12.4568 | -95.1046
12.4571 | -95.1073
12.4574 | -95.1093
12.4578 | -95.1104
12.4581 | -95.1116
12.4585 | -95.113
12.4588 | -95.1142
12.4591 | -95.1156
12.4595 | -95.1171
12.4598 | -95.1183
12.4602 | -95.1198
12.4605 | -95.1211
12.4608 | -95.1222
12.4612 | -95.1236
12.4615 | -95.1249
12.4619 | -95.1261
12.4622 | -95.1274
12.4625 | -95.1287
12.4629 | -95.1302
12.4632 | -95.1318
12.4636 | -95.133
12.4639 | -95.1343
12.4642 | -95.1355
12.4646 | -95.1367
12.4649 | -95.1381
12.4653 | -95.1394
12.4656 | -95.1408
12.466 | -95.142
12.4663 | -95.1438
12.4666 | -95.1456
12.467 | -95.1469
12.4673 | -95.1483
12.4677 | -95.15
12.468 | -95.1515
12.4683 | -95.1529
12.4687 | -95.1541
12.469 | -95.1554
12.4694 | -95.1568
12.4697 | -95.1584
12.47 | -95.1596
12.4704 | -95.1607
12.4707 | -95.1616
12.4711 | -95.1623
12.4714 | -95.1637
12.4717 | -95.1652
12.4721 | -95.1667
12.4724 | -95.1682
12.4728 | -95.1696
12.4731 | -95.1709
12.4734 | -95.1724
12.4738 | -95.1743
12.4741 | -95.1758
12.4745 | -95.177
12.4748 | -95.1784
12.4751 | -95.1801
12.4755 | -95.1814
12.4758 | -95.1827
12.4762 | -95.1846
12.4765 | -95.1862
12.4768 | -95.1877
12.4772 | -95.189
12.4775 | -95.1903
12.4779 | -95.1918
12.4782 | -95.1931
12.4785 | -95.1944
12.4789 | -95.1958
12.4792 | -95.1967
12.4796 | -95.1979
12.4799 | -95.1991
12.4803 | -95.2004
12.4806 | -95.2017
12.4809 | -95.2031
12.4813 | -95.2046
12.4816 | -95.2061
12.482 | -95.2075
12.4823 | -95.2091
12.4826 | -95.2108
12.483 | -95.2123
12.4833 | -95.2136
12.4837 | -95.2152
12.484 | -95.2166
12.4843 | -95.2184
12.4847 | -95.2198
12.485 | -95.2208
12.4854 | -95.2224
12.4857 | -95.2241
12.486 | -95.2251
12.4864 | -95.2259
12.4867 | -95.227
12.4871 | -95.2283
12.4874 | -95.2298
12.4877 | -95.2314
12.4881 | -95.2332
12.4884 | -95.2351
12.4888 | -95.2368
12.4891 | -95.2388
12.4894 | -95.2407
12.4898 | -95.2421
12.4901 | -95.2434
12.4905 | -95.2448
12.4908 | -95.2463
12.4911 | -95.2475
12.4915 | -95.2488
12.4918 | -95.2501
12.4922 | -95.2513
12.4925 | -95.253
12.4929 | -95.2545
12.4932 | -95.2569
12.4935 | -95.2589
12.4939 | -95.2601
12.4942 | -95.2617
12.4946 | -95.264
12.4949 | -95.2662
12.4952 | -95.2682
12.4956 | -95.2698
12.4959 | -95.2709
12.4963 | -95.2723
12.4966 | -95.2735
12.4969 | -95.2745
12.4973 | -95.2758
12.4976 | -95.2774
12.498 | -95.279
12.4983 | -95.2805
12.4986 | -95.2819
12.499 | -95.2836
12.4993 | -95.285
12.4997 | -95.2863
12.5 | -95.2879
12.5003 | -95.2894
12.5007 | -95.2911
12.501 | -95.2936
12.5014 | -95.2966
12.5017 | -95.2991
12.502 | -95.3005
12.5024 | -95.3019
12.5027 | -95.3033
12.5031 | -95.3049
12.5034 | -95.3064
12.5037 | -95.3078
12.5041 | -95.3092
12.5044 | -95.3106
12.5048 | -95.3119
12.5051 | -95.3132
12.5054 | -95.3146
12.5058 | -95.3164
12.5061 | -95.3178
12.5065 | -95.319
12.5068 | -95.3203
12.5071 | -95.3219
12.5075 | -95.3234
12.5078 | -95.325
12.5082 | -95.3265
12.5085 | -95.3278
12.5089 | -95.3294
12.5092 | -95.3311
12.5095 | -95.3326
12.5099 | -95.3339
12.5102 | -95.3355
12.5106 | -95.3369
12.5109 | -95.3384
12.5112 | -95.3396
12.5116 | -95.3404
12.5119 | -95.341
12.5123 | -95.342
12.5126 | -95.3435
12.5129 | -95.345
12.5133 | -95.3466
12.5136 | -95.348
12.514 | -95.3498
12.5143 | -95.3515
12.5146 | -95.3535
12.515 | -95.3549
12.5153 | -95.3563
12.5157 | -95.3582
12.516 | -95.3598
12.5163 | -95.3613
12.5167 | -95.3628
12.517 | -95.3639
12.5174 | -95.3649
12.5177 | -95.3659
12.518 | -95.3664
12.5184 | -95.3681
12.5187 | -95.3699
12.5191 | -95.3713
12.5194 | -95.3721
12.5197 | -95.3737
12.5201 | -95.3753
12.5204 | -95.3767
12.5208 | -95.3784
12.5211 | -95.3798
12.5215 | -95.3815
12.5218 | -95.3831
12.5221 | -95.3845
12.5225 | -95.3859
12.5228 | -95.3872
12.5232 | -95.3887
12.5235 | -95.3897
12.5238 | -95.3904
12.5242 | -95.3919
12.5245 | -95.3934
12.5249 | -95.3949
12.5252 | -95.3965
12.5255 | -95.3982
12.5259 | -95.3997
12.5262 | -95.4013
12.5266 | -95.4026
12.5269 | -95.4039
12.5272 | -95.4051
12.5276 | -95.4066
12.5279 | -95.4084
12.5283 | -95.4101
12.5286 | -95.4118
12.5289 | -95.4134
12.5293 | -95.4148
12.5296 | -95.4165
12.53 | -95.4183
12.5303 | -95.4198
12.5306 | -95.4212
12.531 | -95.4228
12.5313 | -95.4244
12.5317 | -95.4255
12.532 | -95.4262
12.5323 | -95.4272
12.5327 | -95.4287
12.533 | -95.4303
12.5334 | -95.4317
12.5337 | -95.4331
12.534 | -95.4346
12.5344 | -95.436
12.5347 | -95.4376
12.5351 | -95.4384
12.5354 | -95.4397
12.5358 | -95.4412
12.5361 | -95.4426
12.5364 | -95.4433
12.5368 | -95.4441
12.5371 | -95.4454
12.5375 | -95.4469
12.5378 | -95.4481
12.5381 | -95.449
12.5385 | -95.4505
12.5388 | -95.4521
12.5392 | -95.4536
12.5395 | -95.4555
12.5398 | -95.4569
12.5402 | -95.4581
12.5405 | -95.4597
12.5409 | -95.4615
12.5412 | -95.463
12.5415 | -95.4646
12.5419 | -95.4663
12.5422 | -95.4678
12.5426 | -95.469
12.5429 | -95.4702
12.5432 | -95.4715
12.5436 | -95.4728
12.5439 | -95.4741
12.5443 | -95.4756
12.5446 | -95.477
12.5449 | -95.4778
12.5453 | -95.4788
12.5456 | -95.4803
12.546 | -95.4815
12.5463 | -95.4824
12.5466 | -95.4836
12.547 | -95.485
12.5473 | -95.486
12.5477 | -95.4874
12.548 | -95.4889
12.5483 | -95.4903
12.5487 | -95.4912
12.549 | -95.4921
12.5494 | -95.4936
12.5497 | -95.495
12.5501 | -95.4963
12.5504 | -95.4974
12.5507 | -95.4976
12.5511 | -95.4989
12.5514 | -95.4997
12.5518 | -95.4998
12.5521 | -95.4997
12.5524 | -95.5007
12.5528 | -95.5023
12.5531 | -95.5037
12.5535 | -95.5046
12.5538 | -95.5054
12.5541 | -95.5064
12.5545 | -95.5074
12.5548 | -95.5088
12.5552 | -95.5096
12.5555 | -95.5112
12.5558 | -95.5131
12.5562 | -95.5147
12.5565 | -95.5158
12.5569 | -95.5168
12.5572 | -95.5179
12.5575 | -95.5192
12.5579 | -95.5208
12.5582 | -95.522
12.5586 | -95.5233
12.5589 | -95.5248
12.5592 | -95.5263
12.5596 | -95.5273
12.5599 | -95.5284
12.5603 | -95.53
12.5606 | -95.5308
12.5609 | -95.5322
12.5613 | -95.5342
12.5616 | -95.5356
12.562 | -95.5372
12.5623 | -95.5387
12.5626 | -95.5403
12.563 | -95.5417
12.5633 | -95.5431
12.5637 | -95.5444
12.564 | -95.5457
12.5644 | -95.5471
12.5647 | -95.5482
12.565 | -95.5496
12.5654 | -95.5509
12.5657 | -95.5524
12.5661 | -95.5534
12.5664 | -95.5542
12.5667 | -95.5555
12.5671 | -95.5566
12.5674 | -95.5578
12.5678 | -95.559
12.5681 | -95.5599
12.5684 | -95.5614
12.5688 | -95.563
12.5691 | -95.5644
12.5695 | -95.5654
12.5698 | -95.5666
12.5701 | -95.568
12.5705 | -95.5693
12.5708 | -95.5707
12.5712 | -95.5722
12.5715 | -95.5736
12.5718 | -95.5747
12.5722 | -95.5762
12.5725 | -95.5775
12.5729 | -95.5785
12.5732 | -95.5799
12.5735 | -95.5813
12.5739 | -95.5831
12.5742 | -95.5843
12.5746 | -95.5836
12.5749 | -95.585
12.5752 | -95.5855
12.5756 | -95.5864
12.5759 | -95.5877
12.5763 | -95.5887
12.5766 | -95.5897
12.577 | -95.591
12.5773 | -95.5922
12.5776 | -95.5937
12.578 | -95.5952
12.5783 | -95.5959
12.5787 | -95.5971
12.579 | -95.5985
12.5793 | -95.5998
12.5797 | -95.6009
12.58 | -95.6023
12.5804 | -95.6038
12.5807 | -95.6053
12.581 | -95.6067
12.5814 | -95.6078
12.5817 | -95.609
12.5821 | -95.6104
12.5824 | -95.6119
12.5827 | -95.6135
12.5831 | -95.6147
12.5834 | -95.6158
12.5838 | -95.617
12.5841 | -95.6178
12.5844 | -95.6196
12.5848 | -95.6208
12.5851 | -95.622
12.5855 | -95.6233
12.5858 | -95.6245
12.5861 | -95.6257
12.5865 | -95.6272
12.5868 | -95.6287
12.5872 | -95.63
12.5875 | -95.6314
12.5878 | -95.6326
12.5882 | -95.6342
12.5885 | -95.6356
12.5889 | -95.6371
12.5892 | -95.6384
12.5895 | -95.6395
12.5899 | -95.6408
12.5902 | -95.642
12.5906 | -95.6433
12.5909 | -95.6447
12.5913 | -95.6453
12.5916 | -95.6462
12.5919 | -95.6473
12.5923 | -95.6482
12.5926 | -95.6494
12.593 | -95.6507
12.5933 | -95.6518
12.5936 | -95.6532
12.594 | -95.6541
12.5943 | -95.6552
12.5947 | -95.656
12.595 | -95.6574
12.5953 | -95.6585
12.5957 | -95.6598
12.596 | -95.661
12.5964 | -95.6622
12.5967 | -95.6635
12.597 | -95.6648
12.5974 | -95.6662
12.5977 | -95.6674
12.5981 | -95.6684
12.5984 | -95.6698
12.5987 | -95.6707
12.5991 | -95.6717
12.5994 | -95.6729
12.5998 | -95.6738
12.6001 | -95.6748
12.6004 | -95.6762
12.6008 | -95.6776
12.6011 | -95.6789
12.6015 | -95.6803
12.6018 | -95.6816
12.6021 | -95.682
12.6025 | -95.6824
12.6028 | -95.6831
12.6032 | -95.6837
12.6035 | -95.6848
12.6038 | -95.6855
12.6042 | -95.6856
12.6045 | -95.6865
12.6049 | -95.688
12.6052 | -95.6898
12.6056 | -95.6909
12.6059 | -95.692
12.6062 | -95.6929
12.6066 | -95.6935
12.6069 | -95.694
12.6073 | -95.6951
12.6076 | -95.6963
12.6079 | -95.6971
12.6083 | -95.6972
12.6086 | -95.6978
12.609 | -95.699
12.6093 | -95.6997
12.6096 | -95.6997
12.61 | -95.6997
12.6103 | -95.7005
12.6107 | -95.7013
12.611 | -95.7019
12.6113 | -95.7029
12.6117 | -95.7036
12.612 | -95.7041
12.6124 | -95.7049
12.6127 | -95.706
12.613 | -95.7072
12.6134 | -95.7086
12.6137 | -95.7094
12.6141 | -95.7103
12.6144 | -95.7106
12.6147 | -95.71
12.6151 | -95.7101
12.6154 | -95.7114
12.6158 | -95.7124
12.6161 | -95.7125
12.6164 | -95.7132
12.6168 | -95.7141
12.6171 | -95.7144
12.6175 | -95.7148
12.6178 | -95.7158
12.6181 | -95.717
12.6185 | -95.7184
12.6188 | -95.7194
12.6192 | -95.7205
12.6195 | -95.7211
12.6199 | -95.7214
12.6202 | -95.7226
12.6205 | -95.7234
12.6209 | -95.7243
12.6212 | -95.7247
12.6216 | -95.725
12.6219 | -95.7261
12.6222 | -95.7278
12.6226 | -95.7287
12.6229 | -95.7296
12.6233 | -95.7316
12.6236 | -95.7318
12.6239 | -95.7324
12.6243 | -95.7324
12.6246 | -95.7331
12.625 | -95.7344
12.6253 | -95.7356
12.6256 | -95.7365
12.626 | -95.7374
12.6263 | -95.7379
12.6267 | -95.7388
12.627 | -95.7397
12.6273 | -95.7404
12.6277 | -95.7406
12.628 | -95.7409
12.6284 | -95.7416
12.6287 | -95.7426
12.629 | -95.7438
12.6294 | -95.7447
12.6297 | -95.7455
12.6301 | -95.7462
12.6304 | -95.7466
12.6307 | -95.7474
12.6311 | -95.7482
12.6314 | -95.7486
12.6318 | -95.7495
12.6321 | -95.7501
12.6325 | -95.7506
12.6328 | -95.7519
12.6331 | -95.7531
12.6335 | -95.7528
12.6338 | -95.7533
12.6342 | -95.7542
12.6345 | -95.7546
12.6348 | -95.7546
12.6352 | -95.7544
12.6355 | -95.7537
12.6359 | -95.7533
12.6362 | -95.7535
12.6365 | -95.7544
12.6369 | -95.7548
12.6372 | -95.7546
12.6376 | -95.7547
12.6379 | -95.7555
12.6382 | -95.7562
12.6386 | -95.7569
12.6389 | -95.7575
12.6393 | -95.7584
12.6396 | -95.7587
12.6399 | -95.7588
12.6403 | -95.76
12.6406 | -95.7597
12.641 | -95.7594
12.6413 | -95.7601
12.6416 | -95.761
12.642 | -95.7617
12.6423 | -95.762
12.6427 | -95.7622
12.643 | -95.7629
12.6433 | -95.7628
12.6437 | -95.7634
12.644 | -95.764
12.6444 | -95.7644
12.6447 | -95.7648
12.645 | -95.7655
12.6454 | -95.7661
12.6457 | -95.7664
12.6461 | -95.7664
12.6464 | -95.7668
12.6468 | -95.7673
12.6471 | -95.7677
12.6474 | -95.7684
12.6478 | -95.769
12.6481 | -95.7695
12.6485 | -95.7696
12.6488 | -95.7696
12.6491 | -95.7703
12.6495 | -95.7708
12.6498 | -95.7712
12.6502 | -95.7709
12.6505 | -95.7708
12.6508 | -95.7715
12.6512 | -95.7718
12.6515 | -95.7718
12.6519 | -95.7719
12.6522 | -95.7724
12.6525 | -95.773
12.6529 | -95.7722
12.6532 | -95.7724
12.6536 | -95.7713
12.6539 | -95.7707
12.6542 | -95.7711
12.6546 | -95.772
12.6549 | -95.7726
12.6553 | -95.7719
12.6556 | -95.7712
12.6559 | -95.7712
12.6563 | -95.7715
12.6566 | -95.7718
12.657 | -95.772
12.6573 | -95.772
12.6576 | -95.7726
12.658 | -95.7731
12.6583 | -95.7737
12.6587 | -95.7737
12.659 | -95.7739
12.6593 | -95.7744
12.6597 | -95.7749
12.66 | -95.7752
12.6604 | -95.7755
12.6607 | -95.776
12.6611 | -95.776
12.6614 | -95.7752
12.6617 | -95.7755
12.6621 | -95.7756
12.6624 | -95.7758
12.6628 | -95.7746
12.6631 | -95.7753
12.6634 | -95.7759
12.6638 | -95.7763
12.6641 | -95.7767
12.6645 | -95.7771
12.6648 | -95.7775
12.6651 | -95.7776
12.6655 | -95.7773
12.6658 | -95.7774
12.6662 | -95.7778
12.6665 | -95.7783
12.6668 | -95.7779
12.6672 | -95.7777
12.6675 | -95.7781
12.6679 | -95.778
12.6682 | -95.7778
12.6685 | -95.7775
12.6689 | -95.7777
12.6692 | -95.7779
12.6696 | -95.7775
12.6699 | -95.778
12.6702 | -95.7777
12.6706 | -95.7773
12.6709 | -95.7776
12.6713 | -95.7772
12.6716 | -95.7764
12.6719 | -95.7765
12.6723 | -95.7761
12.6726 | -95.7749
12.673 | -95.7749
12.6733 | -95.7755
12.6736 | -95.7753
12.674 | -95.775
12.6743 | -95.7754
12.6747 | -95.7754
12.675 | -95.7751
12.6754 | -95.7745
12.6757 | -95.7745
12.676 | -95.7747
12.6764 | -95.7744
12.6767 | -95.774
12.6771 | -95.7738
12.6774 | -95.7741
12.6777 | -95.7742
12.6781 | -95.7744
12.6784 | -95.7742
12.6788 | -95.7738
12.6791 | -95.7738
12.6794 | -95.774
12.6798 | -95.774
12.6801 | -95.7741
12.6805 | -95.7741
12.6808 | -95.7736
12.6811 | -95.7732
12.6815 | -95.7731
12.6818 | -95.7734
12.6822 | -95.7735
12.6825 | -95.7729
12.6828 | -95.773
12.6832 | -95.7731
12.6835 | -95.773
12.6839 | -95.773
12.6842 | -95.773
12.6845 | -95.7727
12.6849 | -95.7726
12.6852 | -95.7727
12.6856 | -95.7729
12.6859 | -95.7727
12.6862 | -95.7712
12.6866 | -95.7702
12.6869 | -95.7701
12.6873 | -95.7701
12.6876 | -95.7701
12.6879 | -95.7699
12.6883 | -95.7694
12.6886 | -95.7693
12.689 | -95.7691
12.6893 | -95.769
12.6897 | -95.7688
12.69 | -95.7688
12.6903 | -95.769
12.6907 | -95.769
12.691 | -95.7686
12.6914 | -95.7685
12.6917 | -95.7681
12.692 | -95.7677
12.6924 | -95.7665
12.6927 | -95.7665
12.6931 | -95.767
12.6934 | -95.7673
12.6937 | -95.7673
12.6941 | -95.7674
12.6944 | -95.7671
12.6948 | -95.7671
12.6951 | -95.7669
12.6954 | -95.7667
12.6958 | -95.7668
12.6961 | -95.7669
12.6965 | -95.7669
12.6968 | -95.7669
12.6971 | -95.7666
12.6975 | -95.7668
12.6978 | -95.7666
12.6982 | -95.7664
12.6985 | -95.7666
12.6988 | -95.7668
12.6992 | -95.7667
12.6995 | -95.7662
12.6999 | -95.766
12.7002 | -95.7655
12.7005 | -95.7652
12.7009 | -95.7651
12.7012 | -95.7644
12.7016 | -95.7643
12.7019 | -95.7642
12.7022 | -95.7629
12.7026 | -95.7628
12.7029 | -95.7626
12.7033 | -95.7627
12.7036 | -95.7629
12.704 | -95.7626
12.7043 | -95.7623
12.7046 | -95.7621
12.705 | -95.7616
12.7053 | -95.7615
12.7057 | -95.7615
12.706 | -95.761
12.7063 | -95.7608
12.7067 | -95.7607
12.707 | -95.7604
12.7074 | -95.7598
12.7077 | -95.759
12.708 | -95.7591
12.7084 | -95.759
12.7087 | -95.7587
12.7091 | -95.7587
12.7094 | -95.758
12.7097 | -95.7566
12.7101 | -95.7561
12.7104 | -95.7557
12.7108 | -95.7546
12.7111 | -95.7535
12.7114 | -95.7529
12.7118 | -95.7529
12.7121 | -95.7503
12.7125 | -95.7497
12.7128 | -95.7495
12.7131 | -95.7492
12.7135 | -95.7494
12.7138 | -95.7502
12.7142 | -95.75
12.7145 | -95.7495
12.7148 | -95.7491
12.7152 | -95.7487
12.7155 | -95.7484
12.7159 | -95.7479
12.7162 | -95.7474
12.7165 | -95.7466
12.7169 | -95.7458
12.7172 | -95.7452
12.7176 | -95.744
12.7179 | -95.7425
12.7183 | -95.7419
12.7186 | -95.7415
12.7189 | -95.7412
12.7193 | -95.7411
12.7196 | -95.7403
12.72 | -95.7395
12.7203 | -95.739
12.7206 | -95.7383
12.721 | -95.7377
12.7213 | -95.7372
12.7217 | -95.736
12.722 | -95.7355
12.7223 | -95.7345
12.7227 | -95.7332
12.723 | -95.7325
12.7234 | -95.7318
12.7237 | -95.7309
12.724 | -95.7305
12.7244 | -95.7305
12.7247 | -95.7304
12.7251 | -95.7301
12.7254 | -95.7295
12.7257 | -95.729
12.7261 | -95.7286
12.7264 | -95.7288
12.7268 | -95.7284
12.7271 | -95.7278
12.7274 | -95.7271
12.7278 | -95.7262
12.7281 | -95.7257
12.7285 | -95.7247
12.7288 | -95.7229
12.7291 | -95.7219
12.7295 | -95.7215
12.7298 | -95.721
12.7302 | -95.7202
12.7305 | -95.7194
12.7309 | -95.7191
12.7312 | -95.7196
12.7315 | -95.718
12.7319 | -95.7178
12.7322 | -95.7168
12.7326 | -95.7162
12.7329 | -95.7156
12.7332 | -95.7153
12.7336 | -95.7146
12.7339 | -95.713
12.7343 | -95.7127
12.7346 | -95.7119
12.7349 | -95.7106
12.7353 | -95.7105
12.7356 | -95.7098
12.736 | -95.709
12.7363 | -95.708
12.7366 | -95.7072
12.737 | -95.7066
12.7373 | -95.706
12.7377 | -95.7062
12.738 | -95.7056
12.7383 | -95.7047
12.7387 | -95.7043
12.739 | -95.7035
12.7394 | -95.7023
12.7397 | -95.7008
12.74 | -95.6998
12.7404 | -95.6988
12.7407 | -95.698
12.7411 | -95.6971
12.7414 | -95.695
12.7417 | -95.6946
12.7421 | -95.6938
12.7424 | -95.6932
12.7428 | -95.6926
12.7431 | -95.6919
12.7434 | -95.6906
12.7438 | -95.688
12.7441 | -95.6872
12.7445 | -95.6867
12.7448 | -95.6859
12.7452 | -95.6849
12.7455 | -95.684
12.7458 | -95.683
12.7462 | -95.6821
12.7465 | -95.6815
12.7469 | -95.6808
12.7472 | -95.6802
12.7475 | -95.6794
12.7479 | -95.6785
12.7482 | -95.6775
12.7486 | -95.677
12.7489 | -95.6771
12.7492 | -95.6768
12.7496 | -95.6755
12.7499 | -95.6743
12.7503 | -95.673
12.7506 | -95.6719
12.7509 | -95.671
12.7513 | -95.669
12.7516 | -95.6678
12.752 | -95.6671
12.7523 | -95.6665
12.7526 | -95.6654
12.753 | -95.6642
12.7533 | -95.6635
12.7537 | -95.6626
12.754 | -95.6613
12.7543 | -95.6596
12.7547 | -95.6583
12.755 | -95.6576
12.7554 | -95.658
12.7557 | -95.6584
12.756 | -95.658
12.7564 | -95.6568
12.7567 | -95.6553
12.7571 | -95.6542
12.7574 | -95.6535
12.7577 | -95.6526
12.7581 | -95.6513
12.7584 | -95.6501
12.7588 | -95.6486
12.7591 | -95.6472
12.7595 | -95.646
12.7598 | -95.6454
12.7601 | -95.6445
12.7605 | -95.6434
12.7608 | -95.6423
12.7612 | -95.6404
12.7615 | -95.6389
12.7618 | -95.6377
12.7622 | -95.6371
12.7625 | -95.6368
12.7629 | -95.6359
12.7632 | -95.635
12.7635 | -95.6363
12.7639 | -95.638
12.7642 | -95.6501
12.7646 | -95.6675
12.7649 | -95.6776
12.7652 | -95.6796
12.7656 | -95.6793
12.7659 | -95.6776
12.7663 | -95.6763
12.7666 | -95.675
12.7669 | -95.6732
12.7673 | -95.6717
12.7676 | -95.6708
12.768 | -95.6697
12.7683 | -95.6685
12.7686 | -95.6671
12.769 | -95.6656
12.7693 | -95.6642
12.7697 | -95.6629
12.77 | -95.6614
12.7703 | -95.6602
12.7707 | -95.6586
12.771 | -95.6565
12.7714 | -95.6549
12.7717 | -95.6537
12.7721 | -95.6526
12.7724 | -95.6513
12.7727 | -95.6495
12.7731 | -95.648
12.7734 | -95.6464
12.7738 | -95.6449
12.7741 | -95.6438
12.7744 | -95.6427
12.7748 | -95.641
12.7751 | -95.6391
12.7755 | -95.6376
12.7758 | -95.6359
12.7761 | -95.6343
12.7765 | -95.6322
12.7768 | -95.6305
12.7772 | -95.6285
12.7775 | -95.6264
12.7778 | -95.6249
12.7782 | -95.6236
12.7785 | -95.6226
12.7789 | -95.6211
12.7792 | -95.6195
12.7795 | -95.6178
12.7799 | -95.6163
12.7802 | -95.6146
12.7806 | -95.6129
12.7809 | -95.6104
12.7812 | -95.609
12.7816 | -95.6076
12.7819 | -95.6063
12.7823 | -95.6047
12.7826 | -95.6026
12.7829 | -95.6009
12.7833 | -95.599
12.7836 | -95.5971
12.784 | -95.5954
12.7843 | -95.5942
12.7846 | -95.5931
12.785 | -95.591
12.7853 | -95.5892
12.7857 | -95.5878
12.786 | -95.5863
12.7864 | -95.5847
12.7867 | -95.5829
12.787 | -95.5813
12.7874 | -95.5798
12.7877 | -95.5782
12.7881 | -95.5767
12.7884 | -95.5753
12.7887 | -95.5735
12.7891 | -95.5719
12.7894 | -95.57
12.7898 | -95.5684
12.7901 | -95.567
12.7904 | -95.5649
12.7908 | -95.5629
12.7911 | -95.5606
12.7915 | -95.5588
12.7918 | -95.5573
12.7921 | -95.5561
12.7925 | -95.5547
12.7928 | -95.553
12.7932 | -95.551
12.7935 | -95.5493
12.7938 | -95.548
12.7942 | -95.5466
12.7945 | -95.545
12.7949 | -95.5434
12.7952 | -95.5417
12.7955 | -95.5399
12.7959 | -95.5378
12.7962 | -95.5356
12.7966 | -95.534
12.7969 | -95.5323
12.7972 | -95.5307
12.7976 | -95.5292
12.7979 | -95.5272
12.7983 | -95.5253
12.7986 | -95.5233
12.7989 | -95.5204
12.7993 | -95.5173
12.7996 | -95.5149
12.8 | -95.513
12.8003 | -95.5099
12.8007 | -95.5083
12.801 | -95.5066
12.8013 | -95.5048
12.8017 | -95.5021
12.802 | -95.4993
12.8024 | -95.4971
12.8027 | -95.4949
12.803 | -95.4925
12.8034 | -95.4903
12.8037 | -95.4885
12.8041 | -95.4867
12.8044 | -95.4853
12.8047 | -95.4835
12.8051 | -95.4814
12.8054 | -95.4796
12.8058 | -95.4782
12.8061 | -95.4765
12.8064 | -95.4742
12.8068 | -95.4718
12.8071 | -95.4702
12.8075 | -95.4686
12.8078 | -95.4666
12.8081 | -95.4645
12.8085 | -95.4624
12.8088 | -95.4605
12.8092 | -95.4588
12.8095 | -95.457
12.8098 | -95.4551
12.8102 | -95.4518
12.8105 | -95.4498
12.8109 | -95.4479
12.8112 | -95.4459
12.8115 | -95.4437
12.8119 | -95.4417
12.8122 | -95.4399
12.8126 | -95.4377
12.8129 | -95.4357
12.8132 | -95.4338
12.8136 | -95.4319
12.8139 | -95.4299
12.8143 | -95.428
12.8146 | -95.4258
12.815 | -95.4236
12.8153 | -95.4218
12.8156 | -95.4205
12.816 | -95.4192
12.8163 | -95.4174
12.8167 | -95.4155
12.817 | -95.4135
12.8173 | -95.411
12.8177 | -95.4089
12.818 | -95.4071
12.8184 | -95.4051
12.8187 | -95.4032
12.819 | -95.4014
12.8194 | -95.3996
12.8197 | -95.3978
12.8201 | -95.3952
12.8204 | -95.3931
12.8207 | -95.3911
12.8211 | -95.3887
12.8214 | -95.3866
12.8218 | -95.3848
12.8221 | -95.3829
12.8224 | -95.381
12.8228 | -95.3794
12.8231 | -95.3776
12.8235 | -95.3759
12.8238 | -95.3741
12.8241 | -95.3723
12.8245 | -95.3705
12.8248 | -95.3685
12.8252 | -95.3663
12.8255 | -95.3642
12.8258 | -95.3625
12.8262 | -95.3612
12.8265 | -95.3598
12.8269 | -95.358
12.8272 | -95.356
12.8276 | -95.3538
12.8279 | -95.3516
12.8282 | -95.3496
12.8286 | -95.3477
12.8289 | -95.3458
12.8293 | -95.344
12.8296 | -95.3421
12.8299 | -95.3396
12.8303 | -95.3377
12.8306 | -95.3358
12.831 | -95.3337
12.8313 | -95.3316
12.8316 | -95.3296
12.832 | -95.327
12.8323 | -95.3247
12.8327 | -95.3221
12.833 | -95.3196
12.8333 | -95.3175
12.8337 | -95.3157
12.834 | -95.3139
12.8344 | -95.3121
12.8347 | -95.31
12.835 | -95.3077
12.8354 | -95.3057
12.8357 | -95.3035
12.8361 | -95.3018
12.8364 | -95.3004
12.8367 | -95.2988
12.8371 | -95.2968
12.8374 | -95.295
12.8378 | -95.2933
12.8381 | -95.2914
12.8384 | -95.2895
12.8388 | -95.2878
12.8391 | -95.2859
12.8395 | -95.2842
12.8398 | -95.2815
12.8401 | -95.279
12.8405 | -95.277
12.8408 | -95.2752
12.8412 | -95.2734
12.8415 | -95.2716
12.8419 | -95.2696
12.8422 | -95.2678
12.8425 | -95.2658
12.8429 | -95.2641
12.8432 | -95.2622
12.8436 | -95.2601
12.8439 | -95.2579
12.8442 | -95.2557
12.8446 | -95.2535
12.8449 | -95.2514
12.8453 | -95.2499
12.8456 | -95.2478
12.8459 | -95.2458
12.8463 | -95.2441
12.8466 | -95.2422
12.847 | -95.2402
12.8473 | -95.238
12.8476 | -95.236
12.848 | -95.2341
12.8483 | -95.2321
12.8487 | -95.2299
12.849 | -95.228
12.8493 | -95.2261
12.8497 | -95.2239
12.85 | -95.2221
12.8504 | -95.2203
12.8507 | -95.2184
12.851 | -95.2164
12.8514 | -95.2145
12.8517 | -95.2127
12.8521 | -95.2106
12.8524 | -95.208
12.8527 | -95.2057
12.8531 | -95.204
12.8534 | -95.202
12.8538 | -95.2003
12.8541 | -95.1984
12.8544 | -95.1962
12.8548 | -95.1935
12.8551 | -95.1911
12.8555 | -95.1894
12.8558 | -95.1876
12.8562 | -95.1858
12.8565 | -95.184
12.8568 | -95.1821
12.8572 | -95.1799
12.8575 | -95.1781
12.8579 | -95.1763
12.8582 | -95.1742
12.8585 | -95.1723
12.8589 | -95.1703
12.8592 | -95.1683
12.8596 | -95.1664
12.8599 | -95.1644
12.8602 | -95.1626
12.8606 | -95.1605
12.8609 | -95.1582
12.8613 | -95.1563
12.8616 | -95.1543
12.8619 | -95.1523
12.8623 | -95.1505
12.8626 | -95.1487
12.863 | -95.1469
12.8633 | -95.1448
12.8636 | -95.1427
12.864 | -95.14
12.8643 | -95.1373
12.8647 | -95.1343
12.865 | -95.1317
12.8653 | -95.1297
12.8657 | -95.1277
12.866 | -95.1256
12.8664 | -95.1235
12.8667 | -95.1216
12.867 | -95.1199
12.8674 | -95.1179
12.8677 | -95.1157
12.8681 | -95.1138
12.8684 | -95.1119
12.8687 | -95.1098
12.8691 | -95.1072
12.8694 | -95.1049
12.8698 | -95.103
12.8701 | -95.1006
12.8705 | -95.0981
12.8708 | -95.0954
12.8711 | -95.0931
12.8715 | -95.0912
12.8718 | -95.089
12.8722 | -95.0863
12.8725 | -95.0841
12.8728 | -95.0819
12.8732 | -95.0797
12.8735 | -95.0777
12.8739 | -95.0756
12.8742 | -95.0741
12.8745 | -95.0726
12.8749 | -95.0706
12.8752 | -95.0689
12.8756 | -95.0672
12.8759 | -95.0651
12.8762 | -95.0632
12.8766 | -95.0614
12.8769 | -95.0594
12.8773 | -95.0574
12.8776 | -95.0551
12.8779 | -95.0526
12.8783 | -95.0504
12.8786 | -95.0482
12.879 | -95.0456
12.8793 | -95.0436
12.8796 | -95.0415
12.88 | -95.0396
12.8803 | -95.0377
12.8807 | -95.0356
12.881 | -95.0335
12.8813 | -95.0313
12.8817 | -95.0291
12.882 | -95.0268
12.8824 | -95.0243
12.8827 | -95.0224
12.883 | -95.0204
12.8834 | -95.0182
12.8837 | -95.0161
12.8841 | -95.0141
12.8844 | -95.0123
12.8848 | -95.0101
12.8851 | -95.0078
12.8854 | -95.0054
12.8858 | -95.0033
12.8861 | -95.0011
12.8865 | -94.9986
12.8868 | -94.9967
12.8871 | -94.9946
12.8875 | -94.9928
12.8878 | -94.9908
12.8882 | -94.989
12.8885 | -94.9872
12.8888 | -94.9852
12.8892 | -94.9831
12.8895 | -94.9808
12.8899 | -94.9786
12.8902 | -94.9766
12.8905 | -94.9745
12.8909 | -94.9723
12.8912 | -94.9701
12.8916 | -94.9679
12.8919 | -94.9657
12.8922 | -94.9632
12.8926 | -94.961
12.8929 | -94.959
12.8933 | -94.9569
12.8936 | -94.9549
12.8939 | -94.9526
12.8943 | -94.9503
12.8946 | -94.9483
12.895 | -94.9463
12.8953 | -94.9443
12.8956 | -94.9423
12.896 | -94.9407
12.8963 | -94.939
12.8967 | -94.9369
12.897 | -94.9351
12.8973 | -94.933
12.8977 | -94.9308
12.898 | -94.9287
12.8984 | -94.9267
12.8987 | -94.9248
12.8991 | -94.9225
12.8994 | -94.9204
12.8997 | -94.9183
12.9001 | -94.9163
12.9004 | -94.9141
12.9008 | -94.9121
12.9011 | -94.9102
12.9014 | -94.908
12.9018 | -94.9059
12.9021 | -94.9039
12.9025 | -94.9017
12.9028 | -94.8998
12.9031 | -94.8977
12.9035 | -94.8958
12.9038 | -94.8941
12.9042 | -94.8923
12.9045 | -94.8904
12.9048 | -94.8883
12.9052 | -94.8863
12.9055 | -94.8845
12.9059 | -94.883
12.9062 | -94.8815
12.9065 | -94.8796
12.9069 | -94.8776
12.9072 | -94.8756
12.9076 | -94.8736
12.9079 | -94.8717
12.9082 | -94.8697
12.9086 | -94.868
12.9089 | -94.8659
12.9093 | -94.864
12.9096 | -94.862
12.9099 | -94.8601
12.9103 | -94.8584
12.9106 | -94.8566
12.911 | -94.8549
12.9113 | -94.8527
12.9116 | -94.8501
12.912 | -94.8477
12.9123 | -94.8457
12.9127 | -94.8438
12.913 | -94.8418
12.9134 | -94.8397
12.9137 | -94.8376
12.914 | -94.8353
12.9144 | -94.8332
12.9147 | -94.8311
12.9151 | -94.8292
12.9154 | -94.8273
12.9157 | -94.8253
12.9161 | -94.8235
12.9164 | -94.8217
12.9168 | -94.82
12.9171 | -94.8185
12.9174 | -94.8167
12.9178 | -94.8148
12.9181 | -94.8132
12.9185 | -94.8114
12.9188 | -94.8096
12.9191 | -94.8077
12.9195 | -94.8058
12.9198 | -94.8038
12.9202 | -94.8017
12.9205 | -94.7995
12.9208 | -94.7974
12.9212 | -94.7951
12.9215 | -94.7931
12.9219 | -94.7911
12.9222 | -94.7889
12.9225 | -94.7862
12.9229 | -94.784
12.9232 | -94.7821
12.9236 | -94.7803
12.9239 | -94.7781
12.9242 | -94.7757
12.9246 | -94.7737
12.9249 | -94.7719
12.9253 | -94.7702
12.9256 | -94.7684
12.926 | -94.7667
12.9263 | -94.7646
12.9266 | -94.762
12.927 | -94.76
12.9273 | -94.7581
12.9277 | -94.7562
12.928 | -94.7545
12.9283 | -94.7528
12.9287 | -94.751
12.929 | -94.7491
12.9294 | -94.7473
12.9297 | -94.7455
12.93 | -94.7437
12.9304 | -94.7418
12.9307 | -94.74
12.9311 | -94.7381
12.9314 | -94.7361
12.9317 | -94.734
12.9321 | -94.7323
12.9324 | -94.7304
12.9328 | -94.7286
12.9331 | -94.7267
12.9334 | -94.7249
12.9338 | -94.7231
12.9341 | -94.7215
12.9345 | -94.7197
12.9348 | -94.7179
12.9351 | -94.716
12.9355 | -94.7142
12.9358 | -94.7124
12.9362 | -94.7104
12.9365 | -94.7085
12.9368 | -94.7066
12.9372 | -94.7047
12.9375 | -94.7027
12.9379 | -94.7006
12.9382 | -94.6985
12.9385 | -94.6969
12.9389 | -94.6953
12.9392 | -94.6936
12.9396 | -94.6918
12.9399 | -94.6899
12.9403 | -94.688
12.9406 | -94.6863
12.9409 | -94.6843
12.9413 | -94.6819
12.9416 | -94.6801
12.942 | -94.6782
12.9423 | -94.6762
12.9426 | -94.6743
12.943 | -94.6726
12.9433 | -94.671
12.9437 | -94.6692
12.944 | -94.6674
12.9443 | -94.6657
12.9447 | -94.6639
12.945 | -94.6624
12.9454 | -94.6608
12.9457 | -94.6592
12.946 | -94.6578
12.9464 | -94.656
12.9467 | -94.6544
12.9471 | -94.653
12.9474 | -94.6513
12.9477 | -94.6499
12.9481 | -94.648
12.9484 | -94.6465
12.9488 | -94.645
12.9491 | -94.6435
12.9494 | -94.6417
12.9498 | -94.6401
12.9501 | -94.6397
12.9505 | -94.6391
12.9508 | -94.6396
12.9511 | -94.6402
12.9515 | -94.6386
12.9518 | -94.6372
12.9522 | -94.6355
12.9525 | -94.6335
12.9528 | -94.6326
12.9532 | -94.6319
12.9535 | -94.6315
12.9539 | -94.6313
12.9542 | -94.6304
12.9546 | -94.6287
12.9549 | -94.6269
12.9552 | -94.6251
12.9556 | -94.6237
12.9559 | -94.6234
12.9563 | -94.6234
12.9566 | -94.624
12.9569 | -94.624
12.9573 | -94.6222
12.9576 | -94.6208
12.958 | -94.619
12.9583 | -94.6172
12.9586 | -94.6159
12.959 | -94.6145
12.9593 | -94.6129
12.9597 | -94.6113
12.96 | -94.6096
12.9603 | -94.6079
12.9607 | -94.6062
12.961 | -94.6048
12.9614 | -94.6034
12.9617 | -94.6023
12.962 | -94.6005
12.9624 | -94.5986
12.9627 | -94.5969
12.9631 | -94.5951
12.9634 | -94.5934
12.9637 | -94.5916
12.9641 | -94.5903
12.9644 | -94.5888
12.9648 | -94.5872
12.9651 | -94.5855
12.9654 | -94.5838
12.9658 | -94.5818
12.9661 | -94.5796
12.9665 | -94.5778
12.9668 | -94.5763
12.9672 | -94.5747
12.9675 | -94.5729
12.9678 | -94.5713
12.9682 | -94.5697
12.9685 | -94.5681
12.9689 | -94.5662
12.9692 | -94.5644
12.9695 | -94.5627
12.9699 | -94.5608
12.9702 | -94.5587
12.9706 | -94.5567
12.9709 | -94.5551
12.9712 | -94.5534
12.9716 | -94.5516
12.9719 | -94.5498
12.9723 | -94.5483
12.9726 | -94.5468
12.9729 | -94.5455
12.9733 | -94.5439
12.9736 | -94.542
12.974 | -94.5402
12.9743 | -94.5386
12.9746 | -94.5371
12.975 | -94.5354
12.9753 | -94.5336
12.9757 | -94.5314
12.976 | -94.5296
12.9763 | -94.5279
12.9767 | -94.5261
12.977 | -94.5244
12.9774 | -94.5232
12.9777 | -94.5214
12.978 | -94.5197
12.9784 | -94.5183
12.9787 | -94.5169
12.9791 | -94.5153
12.9794 | -94.5137
12.9797 | -94.5124
12.9801 | -94.5111
12.9804 | -94.5095
12.9808 | -94.5074
12.9811 | -94.5053
12.9815 | -94.5036
12.9818 | -94.5021
12.9821 | -94.5005
12.9825 | -94.4987
12.9828 | -94.4971
12.9832 | -94.4955
12.9835 | -94.494
12.9838 | -94.4924
12.9842 | -94.491
12.9845 | -94.4894
12.9849 | -94.4879
12.9852 | -94.4863
12.9855 | -94.4848
12.9859 | -94.4833
12.9862 | -94.4817
12.9866 | -94.4794
12.9869 | -94.4783
12.9872 | -94.4767
12.9876 | -94.475
12.9879 | -94.4734
12.9883 | -94.4718
12.9886 | -94.4703
12.9889 | -94.4686
12.9893 | -94.4669
12.9896 | -94.4655
12.99 | -94.4641
12.9903 | -94.4625
12.9906 | -94.461
12.991 | -94.4596
12.9913 | -94.4579
12.9917 | -94.4564
12.992 | -94.455
12.9923 | -94.4536
12.9927 | -94.4522
12.993 | -94.4506
12.9934 | -94.4491
12.9937 | -94.4477
12.994 | -94.446
12.9944 | -94.4446
12.9947 | -94.443
12.9951 | -94.4415
12.9954 | -94.44
12.9958 | -94.4386
12.9961 | -94.4372
12.9964 | -94.4355
12.9968 | -94.4345
12.9971 | -94.4331
12.9975 | -94.4315
12.9978 | -94.4297
12.9981 | -94.4281
12.9985 | -94.4266
12.9988 | -94.4253
12.9992 | -94.424
12.9995 | -94.4228
12.9998 | -94.4214
13.0002 | -94.4197
13.0005 | -94.4182
13.0009 | -94.4169
13.0012 | -94.4154
13.0015 | -94.4138
13.0019 | -94.4121
13.0022 | -94.4107
13.0026 | -94.4095
13.0029 | -94.4082
13.0032 | -94.4068
13.0036 | -94.4056
13.0039 | -94.4041
13.0043 | -94.4026
13.0046 | -94.4013
13.0049 | -94.4002
13.0053 | -94.3989
13.0056 | -94.3976
13.006 | -94.3957
13.0063 | -94.3942
13.0066 | -94.3936
13.007 | -94.3921
13.0073 | -94.3907
13.0077 | -94.3894
13.008 | -94.3879
13.0083 | -94.3864
13.0087 | -94.385
13.009 | -94.3835
13.0094 | -94.3819
13.0097 | -94.3804
13.0101 | -94.379
13.0104 | -94.3777
13.0107 | -94.3763
13.0111 | -94.3751
13.0114 | -94.3737
13.0118 | -94.3724
13.0121 | -94.3711
13.0124 | -94.3696
13.0128 | -94.3683
13.0131 | -94.3668
13.0135 | -94.3652
13.0138 | -94.3633
13.0141 | -94.3619
13.0145 | -94.3609
13.0148 | -94.36
13.0152 | -94.3589
13.0155 | -94.3577
13.0158 | -94.3564
13.0162 | -94.3548
13.0165 | -94.3542
13.0169 | -94.3533
13.0172 | -94.352
13.0175 | -94.3507
13.0179 | -94.3491
13.0182 | -94.3475
13.0186 | -94.3463
13.0189 | -94.3451
13.0192 | -94.3438
13.0196 | -94.3426
13.0199 | -94.3415
13.0203 | -94.3402
13.0206 | -94.339
13.0209 | -94.3378
13.0213 | -94.3365
13.0216 | -94.3349
13.022 | -94.3336
13.0223 | -94.3325
13.0226 | -94.3313
13.023 | -94.33
13.0233 | -94.3288
13.0237 | -94.3277
13.024 | -94.3267
13.0244 | -94.3253
13.0247 | -94.3241
13.025 | -94.3231
13.0254 | -94.3219
13.0257 | -94.3208
13.0261 | -94.3199
13.0264 | -94.3196
13.0267 | -94.3186
13.0271 | -94.3174
13.0274 | -94.3162
13.0278 | -94.315
13.0281 | -94.3139
13.0284 | -94.313
13.0288 | -94.3121
13.0291 | -94.311
13.0295 | -94.3095
13.0298 | -94.3082
13.0301 | -94.3069
13.0305 | -94.3058
13.0308 | -94.3047
13.0312 | -94.3035
13.0315 | -94.3023
13.0318 | -94.3012
13.0322 | -94.2999
13.0325 | -94.2988
13.0329 | -94.2978
13.0332 | -94.2969
13.0335 | -94.2958
13.0339 | -94.2947
13.0342 | -94.2936
13.0346 | -94.2925
13.0349 | -94.2914
13.0352 | -94.2902
13.0356 | -94.289
13.0359 | -94.288
13.0363 | -94.2877
13.0366 | -94.2866
13.0369 | -94.2856
13.0373 | -94.2846
13.0376 | -94.2836
13.038 | -94.2825
13.0383 | -94.2815
13.0387 | -94.2804
13.039 | -94.2792
13.0393 | -94.2781
13.0397 | -94.2769
13.04 | -94.2759
13.0404 | -94.2748
13.0407 | -94.2738
13.041 | -94.2728
13.0414 | -94.2718
13.0417 | -94.271
13.0421 | -94.2701
13.0424 | -94.2693
13.0427 | -94.2683
13.0431 | -94.2673
13.0434 | -94.2662
13.0438 | -94.2651
13.0441 | -94.2642
13.0444 | -94.2634
13.0448 | -94.2625
13.0451 | -94.2616
13.0455 | -94.2606
13.0458 | -94.2594
13.0461 | -94.2593
13.0465 | -94.2583
13.0468 | -94.2574
13.0472 | -94.2565
13.0475 | -94.2556
13.0478 | -94.2547
13.0482 | -94.2537
13.0485 | -94.2526
13.0489 | -94.2514
13.0492 | -94.2503
13.0495 | -94.2495
13.0499 | -94.2489
13.0502 | -94.248
13.0506 | -94.2473
13.0509 | -94.2465
13.0512 | -94.2457
13.0516 | -94.2448
13.0519 | -94.2438
13.0523 | -94.2429
13.0526 | -94.242
13.053 | -94.2411
13.0533 | -94.2401
13.0536 | -94.2391
13.054 | -94.2382
13.0543 | -94.2374
13.0547 | -94.2366
13.055 | -94.2358
13.0553 | -94.235
13.0557 | -94.2352
13.056 | -94.2346
13.0564 | -94.2339
13.0567 | -94.2331
13.057 | -94.2322
13.0574 | -94.231
13.0577 | -94.23
13.0581 | -94.2291
13.0584 | -94.2283
13.0587 | -94.2275
13.0591 | -94.2268
13.0594 | -94.226
13.0598 | -94.2253
13.0601 | -94.2245
13.0604 | -94.2237
13.0608 | -94.2229
13.0611 | -94.2222
13.0615 | -94.2216
13.0618 | -94.2208
13.0621 | -94.22
13.0625 | -94.2192
13.0628 | -94.2185
13.0632 | -94.2178
13.0635 | -94.2171
13.0638 | -94.2163
13.0642 | -94.2155
13.0645 | -94.2149
13.0649 | -94.214
13.0652 | -94.2132
13.0656 | -94.2135
13.0659 | -94.2128
13.0662 | -94.212
13.0666 | -94.2114
13.0669 | -94.211
13.0673 | -94.2103
13.0676 | -94.2095
13.0679 | -94.2091
13.0683 | -94.2086
13.0686 | -94.2081
13.069 | -94.2074
13.0693 | -94.2066
13.0696 | -94.2058
13.07 | -94.2051
13.0703 | -94.2043
13.0707 | -94.2035
13.071 | -94.2029
13.0713 | -94.2022
13.0717 | -94.2015
13.072 | -94.2007
13.0724 | -94.2001
13.0727 | -94.1994
13.073 | -94.1986
13.0734 | -94.1978
13.0737 | -94.1971
13.0741 | -94.1966
13.0744 | -94.1963
13.0747 | -94.1961
13.0751 | -94.1958
13.0754 | -94.1963
13.0758 | -94.1954
13.0761 | -94.1944
13.0764 | -94.1937
13.0768 | -94.193
13.0771 | -94.1922
13.0775 | -94.1914
13.0778 | -94.1906
13.0781 | -94.19
13.0785 | -94.1894
13.0788 | -94.1888
13.0792 | -94.1884
13.0795 | -94.1878
13.0799 | -94.1872
13.0802 | -94.1866
13.0805 | -94.1859
13.0809 | -94.1855
13.0812 | -94.185
13.0816 | -94.1847
13.0819 | -94.1842
13.0822 | -94.1836
13.0826 | -94.183
13.0829 | -94.1823
13.0833 | -94.1818
13.0836 | -94.1812
13.0839 | -94.1806
13.0843 | -94.1801
13.0846 | -94.1796
13.085 | -94.1796
13.0853 | -94.18
13.0856 | -94.1796
13.086 | -94.1788
13.0863 | -94.1786
13.0867 | -94.1779
13.087 | -94.1774
13.0873 | -94.1768
13.0877 | -94.1764
13.088 | -94.176
13.0884 | -94.1753
13.0887 | -94.1747
13.089 | -94.1742
13.0894 | -94.1737
13.0897 | -94.1734
13.0901 | -94.1727
13.0904 | -94.172
13.0907 | -94.1715
13.0911 | -94.1711
13.0914 | -94.171
13.0918 | -94.1706
13.0921 | -94.1702
13.0924 | -94.1698
13.0928 | -94.1693
13.0931 | -94.1688
13.0935 | -94.1684
13.0938 | -94.1681
13.0942 | -94.1677
13.0945 | -94.1675
13.0948 | -94.1672
13.0952 | -94.1679
13.0955 | -94.1675
13.0959 | -94.1671
13.0962 | -94.1666
13.0965 | -94.1662
13.0969 | -94.1659
13.0972 | -94.1655
13.0976 | -94.1651
13.0979 | -94.1648
13.0982 | -94.1645
13.0986 | -94.1645
13.0989 | -94.1642
13.0993 | -94.1639
13.0996 | -94.1635
13.0999 | -94.1631
13.1003 | -94.1628
13.1006 | -94.1625
13.101 | -94.1623
13.1013 | -94.1621
13.1016 | -94.1617
13.102 | -94.1613
13.1023 | -94.1609
13.1027 | -94.1602
13.103 | -94.1595
13.1033 | -94.1593
13.1037 | -94.159
13.104 | -94.1587
13.1044 | -94.1585
13.1047 | -94.1583
13.105 | -94.159
13.1054 | -94.1588
13.1057 | -94.1586
13.1061 | -94.1583
13.1064 | -94.1581
13.1067 | -94.158
13.1071 | -94.1578
13.1074 | -94.1575
13.1078 | -94.1573
13.1081 | -94.1569
13.1085 | -94.1567
13.1088 | -94.1565
13.1091 | -94.1562
13.1095 | -94.1559
13.1098 | -94.1556
13.1102 | -94.1553
13.1105 | -94.1551
13.1108 | -94.1548
13.1112 | -94.1545
13.1115 | -94.1543
13.1119 | -94.1541
13.1122 | -94.1538
13.1125 | -94.1536
13.1129 | -94.1535
13.1132 | -94.1534
13.1136 | -94.1532
13.1139 | -94.1531
13.1142 | -94.1531
13.1146 | -94.1531
13.1149 | -94.1539
13.1153 | -94.1536
13.1156 | -94.1532
13.1159 | -94.1528
13.1163 | -94.1525
13.1166 | -94.1523
13.117 | -94.1523
13.1173 | -94.1522
13.1176 | -94.1521
13.118 | -94.1521
13.1183 | -94.1525
13.1187 | -94.1526
13.119 | -94.1527
13.1193 | -94.1527
13.1197 | -94.1525
13.12 | -94.1522
13.1204 | -94.152
13.1207 | -94.1518
13.1211 | -94.1518
13.1214 | -94.1517
13.1217 | -94.1518
13.1221 | -94.1519
13.1224 | -94.1521
13.1228 | -94.1521
13.1231 | -94.1522
13.1234 | -94.1521
13.1238 | -94.1522
13.1241 | -94.1522
13.1245 | -94.1531
13.1248 | -94.1531
13.1251 | -94.153
13.1255 | -94.153
13.1258 | -94.153
13.1262 | -94.153
13.1265 | -94.1528
13.1268 | -94.1528
13.1272 | -94.153
13.1275 | -94.153
13.1279 | -94.1531
13.1282 | -94.153
13.1285 | -94.153
13.1289 | -94.1532
13.1292 | -94.1532
13.1296 | -94.1534
13.1299 | -94.1535
13.1302 | -94.1535
13.1306 | -94.1536
13.1309 | -94.1534
13.1313 | -94.1533
13.1316 | -94.1534
13.1319 | -94.1534
13.1323 | -94.1537
13.1326 | -94.1543
13.133 | -94.1545
13.1333 | -94.1543
13.1336 | -94.1542
13.134 | -94.1544
13.1343 | -94.1556
13.1347 | -94.156
13.135 | -94.1563
13.1354 | -94.1564
13.1357 | -94.1564
13.136 | -94.1565
13.1364 | -94.1565
13.1367 | -94.1566
13.1371 | -94.1566
13.1374 | -94.1566
13.1377 | -94.1565
13.1381 | -94.1564
13.1384 | -94.1565
13.1388 | -94.1562
13.1391 | -94.1561
13.1394 | -94.1563
13.1398 | -94.1566
13.1401 | -94.1569
13.1405 | -94.1574
13.1408 | -94.1576
13.1411 | -94.1577
13.1415 | -94.1579
13.1418 | -94.1581
13.1422 | -94.1583
13.1425 | -94.1586
13.1428 | -94.1588
13.1432 | -94.1589
13.1435 | -94.1592
13.1439 | -94.1594
13.1442 | -94.1607
13.1445 | -94.1608
13.1449 | -94.1611
13.1452 | -94.1613
13.1456 | -94.1615
13.1459 | -94.1617
13.1462 | -94.162
13.1466 | -94.1623
13.1469 | -94.1627
13.1473 | -94.163
13.1476 | -94.1632
13.1479 | -94.1633
13.1483 | -94.1635
13.1486 | -94.164
13.149 | -94.164
13.1493 | -94.1641
13.1497 | -94.1644
13.15 | -94.1647
13.1503 | -94.165
13.1507 | -94.1653
13.151 | -94.1656
13.1514 | -94.1659
13.1517 | -94.1661
13.152 | -94.1665
13.1524 | -94.1671
13.1527 | -94.1673
13.1531 | -94.1676
13.1534 | -94.1676
13.1537 | -94.1677
13.1541 | -94.1692
13.1544 | -94.1695
13.1548 | -94.1699
13.1551 | -94.1703
13.1554 | -94.1708
13.1558 | -94.1712
13.1561 | -94.1716
13.1565 | -94.1719
13.1568 | -94.1721
13.1571 | -94.1721
13.1575 | -94.1723
13.1578 | -94.1727
13.1582 | -94.1731
13.1585 | -94.1732
13.1588 | -94.1735
13.1592 | -94.1739
13.1595 | -94.1743
13.1599 | -94.1747
13.1602 | -94.175
13.1605 | -94.1754
13.1609 | -94.1755
13.1612 | -94.1759
13.1616 | -94.1763
13.1619 | -94.1768
13.1622 | -94.1775
13.1626 | -94.1777
13.1629 | -94.1778
13.1633 | -94.1779
13.1636 | -94.1787
13.164 | -94.1801
13.1643 | -94.1807
13.1646 | -94.1812
13.165 | -94.1816
13.1653 | -94.182
13.1657 | -94.1825
13.166 | -94.1828
13.1663 | -94.1833
13.1667 | -94.1837
13.167 | -94.1841
13.1674 | -94.1844
13.1677 | -94.1848
13.168 | -94.1853
13.1684 | -94.1858
13.1687 | -94.1864
13.1691 | -94.187
13.1694 | -94.1876
13.1697 | -94.188
13.1701 | -94.1886
13.1704 | -94.1892
13.1708 | -94.1897
13.1711 | -94.1903
13.1714 | -94.1912
13.1718 | -94.1917
13.1721 | -94.1921
13.1725 | -94.1926
13.1728 | -94.193
13.1731 | -94.1935
13.1735 | -94.194
13.1738 | -94.1955
13.1742 | -94.1959
13.1745 | -94.1962
13.1748 | -94.1966
13.1752 | -94.197
13.1755 | -94.1975
13.1759 | -94.1982
13.1762 | -94.199
13.1766 | -94.1996
13.1769 | -94.2001
13.1772 | -94.2007
13.1776 | -94.2012
13.1779 | -94.2016
13.1783 | -94.202
13.1786 | -94.2027
13.1789 | -94.2033
13.1793 | -94.2039
13.1796 | -94.2044
13.18 | -94.2049
13.1803 | -94.2055
13.1806 | -94.2063
13.181 | -94.207
13.1813 | -94.2077
13.1817 | -94.2082
13.182 | -94.2089
13.1823 | -94.2095
13.1827 | -94.2101
13.183 | -94.2112
13.1834 | -94.2128
13.1837 | -94.2137
13.184 | -94.2143
13.1844 | -94.2148
13.1847 | -94.2154
13.1851 | -94.2162
13.1854 | -94.2169
13.1857 | -94.2176
13.1861 | -94.2182
13.1864 | -94.2189
13.1868 | -94.2197
13.1871 | -94.2205
13.1874 | -94.2211
13.1878 | -94.2212
13.1881 | -94.2217
13.1885 | -94.2226
13.1888 | -94.2232
13.1891 | -94.2238
13.1895 | -94.2246
13.1898 | -94.2254
13.1902 | -94.2262
13.1905 | -94.2269
13.1908 | -94.2277
13.1912 | -94.2284
13.1915 | -94.2291
13.1919 | -94.2298
13.1922 | -94.2307
13.1926 | -94.2316
13.1929 | -94.2324
13.1932 | -94.2337
13.1936 | -94.2345
13.1939 | -94.2351
13.1943 | -94.2359
13.1946 | -94.2367
13.1949 | -94.2374
13.1953 | -94.2383
13.1956 | -94.2391
13.196 | -94.2398
13.1963 | -94.2404
13.1966 | -94.241
13.197 | -94.2416
13.1973 | -94.2424
13.1977 | -94.2434
13.198 | -94.2443